## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 290
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 2.1156 | 0.066414 | 1.9828 | 2.2484 | 1.9163 | 2.3148 | 3.139% | 6.279% | 9.418% | 2.0098 | 2.2214
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.9323 | 0.090981 | 1.7503 | 2.1143 | 1.6594 | 2.2052 | 4.708% | 9.417% | 14.13% | 1.8357 | 2.0289
Peroxide Fusion ICP
Sb, wt.% | 0.79545 | 0.030257 | 0.73494 | 0.85597 | 0.70468 | 0.88622 | 3.804% | 7.607% | 11.41% | 0.75568 | 0.83523
4-Acid Digestion
Ag, ppm | 0.21349 | 0.040128 | 0.13324 | 0.29375 | 0.093108 | 0.33388 | 18.8% | 37.59% | 56.39% | 0.20282 | 0.22417
Al, wt.% | 7.2094 | 0.27234 | 6.6647 | 7.754 | 6.3924 | 8.0264 | 3.778% | 7.555% | 11.33% | 6.8489 | 7.5698
As, ppm | 313.437 | 19.102 | 275.2331 | 351.6409 | 256.1311 | 370.7429 | 6.094% | 12.19% | 18.28% | 297.7652 | 329.1089
Ba, ppm | 738.9002 | 32.2298 | 674.4406 | 803.3599 | 642.2107 | 835.5897 | 4.362% | 8.724% | 13.09% | 701.9552 | 775.8452
Be, ppm | 2.5085 | 0.19655 | 2.1154 | 2.9016 | 1.9188 | 3.0982 | 7.835% | 15.67% | 23.51% | 2.3831 | 2.6339
Bi, ppm | 0.36738 | 0.03128 | 0.30482 | 0.42994 | 0.27354 | 0.46122 | 8.514% | 17.03% | 25.54% | 0.34901 | 0.38575
Ca, wt.% | 0.52455 | 0.030756 | 0.46303 | 0.58606 | 0.43228 | 0.61681 | 5.863% | 11.73% | 17.59% | 0.49832 | 0.55077
Cd, ppm | 0.053506 | 0.012649 | 0.028209 | 0.078804 | 0.015561 | 0.091452 | 23.64% | 47.28% | 70.92% | 0.050831 | 0.056182
Ce, ppm | 78.7458 | 6.3467 | 66.0524 | 91.4391 | 59.7057 | 97.7858 | 8.06% | 16.12% | 24.18% | 74.8085 | 82.683
Co, ppm | 14.1993 | 0.91894 | 12.3615 | 16.0372 | 11.4425 | 16.9562 | 6.472% | 12.94% | 19.42% | 13.4894 | 14.9093
Cr, ppm | 122.7493 | 6.8127 | 109.124 | 136.3747 | 102.3114 | 143.1873 | 5.55% | 11.1% | 16.65% | 116.6119 | 128.8868
Cs, ppm | 10.33 | 0.62089 | 9.0882 | 11.5717 | 8.4673 | 12.1926 | 6.011% | 12.02% | 18.03% | 9.8135 | 10.8465
Cu, ppm | 32.5699 | 1.9399 | 28.69 | 36.4497 | 26.7501 | 38.3896 | 5.956% | 11.91% | 17.87% | 30.9414 | 34.1984
Dy, ppm | 3.0127 | 0.35006 | 2.3125 | 3.7128 | 1.9625 | 4.0629 | 11.62% | 23.24% | 34.86% | 2.862 | 3.1633
Er, ppm | 1.6002 | 0.17804 | 1.2441 | 1.9563 | 1.0661 | 2.1343 | 11.13% | 22.25% | 33.38% | 1.5202 | 1.6802
Eu, ppm | 1.141 | 0.096895 | 0.94721 | 1.3348 | 0.85032 | 1.4317 | 8.492% | 16.98% | 25.48% | 1.084 | 1.1981
Fe, wt.% | 3.7892 | 0.1902 | 3.4088 | 4.1696 | 3.2186 | 4.3598 | 5.019% | 10.04% | 15.06% | 3.5998 | 3.9787
Ga, ppm | 18.6282 | 1.1726 | 16.283 | 20.9734 | 15.1104 | 22.146 | 6.295% | 12.59% | 18.88% | 17.6968 | 19.5596
Gd, ppm | 4.4747 | 0.43575 | 3.6032 | 5.3462 | 3.1674 | 5.7819 | 9.738% | 19.48% | 29.21% | 4.2509 | 4.6984
Ge, ppm | 0.2172 | 0.059718 | 0.097764 | 0.33664 | 0.038046 | 0.39635 | 27.49% | 54.99% | 82.48% | 0.20634 | 0.22806
Hf, ppm | 3.9792 | 0.26097 | 3.4573 | 4.5012 | 3.1963 | 4.7622 | 6.558% | 13.12% | 19.68% | 3.7803 | 4.1782
Ho, ppm | 0.56167 | 0.067121 | 0.42743 | 0.69591 | 0.3603 | 0.76303 | 11.95% | 23.9% | 35.85% | 0.53358 | 0.58975
In, ppm | 0.064435 | 0.0068814 | 0.050672 | 0.078198 | 0.043791 | 0.085079 | 10.68% | 21.36% | 32.04% | 0.061213 | 0.067657
K, wt.% | 2.8003 | 0.09941 | 2.6015 | 2.9992 | 2.5021 | 3.0986 | 3.55% | 7.1% | 10.65% | 2.6603 | 2.9404
La, ppm | 39.5905 | 3.3778 | 32.8348 | 46.3462 | 29.457 | 49.7241 | 8.532% | 17.06% | 25.6% | 37.611 | 41.57
Li, ppm | 51.3773 | 2.8458 | 45.6857 | 57.069 | 42.8398 | 59.9148 | 5.539% | 11.08% | 16.62% | 48.8085 | 53.9462
Lu, ppm | 0.23619 | 0.038695 | 0.1588 | 0.31358 | 0.1201 | 0.35228 | 16.38% | 32.77% | 49.15% | 0.22438 | 0.248
Mg, wt.% | 1.388 | 0.072345 | 1.2434 | 1.5327 | 1.171 | 1.6051 | 5.212% | 10.42% | 15.64% | 1.3186 | 1.4575
Mn, wt.% | 0.056358 | 0.002214 | 0.05193 | 0.060786 | 0.049716 | 0.063 | 3.928% | 7.857% | 11.79% | 0.05354 | 0.059176
Mo, ppm | 1.1356 | 0.16886 | 0.79785 | 1.4733 | 0.62899 | 1.6422 | 14.87% | 29.74% | 44.61% | 1.0788 | 1.1924
Na, wt.% | 0.61344 | 0.024735 | 0.56397 | 0.66291 | 0.53924 | 0.68765 | 4.032% | 8.064% | 12.1% | 0.58277 | 0.64411
Nb, ppm | 10.1868 | 2.8738 | 4.4391 | 15.9344 | 1.5652 | 18.8083 | 28.21% | 56.42% | 84.63% | 9.6774 | 10.6961
Nd, ppm | 34.2123 | 2.2574 | 29.6975 | 38.7272 | 27.44 | 40.9846 | 6.598% | 13.2% | 19.79% | 32.5017 | 35.923
Ni, ppm | 72.591 | 3.5151 | 65.5608 | 79.6213 | 62.0457 | 83.1364 | 4.842% | 9.685% | 14.53% | 68.9615 | 76.2206
P, wt.% | 0.05857 | 0.0030789 | 0.052412 | 0.064728 | 0.049333 | 0.067807 | 5.257% | 10.51% | 15.77% | 0.055642 | 0.061499
Pb, ppm | 22.8867 | 1.3821 | 20.1224 | 25.651 | 18.7402 | 27.0331 | 6.039% | 12.08% | 18.12% | 21.7423 | 24.031
Pr, ppm | 9.1335 | 0.59428 | 7.945 | 10.3221 | 7.3507 | 10.9164 | 6.507% | 13.01% | 19.52% | 8.6769 | 9.5902
Rb, ppm | 161.2573 | 10.9884 | 139.2804 | 183.2341 | 128.292 | 194.2225 | 6.814% | 13.63% | 20.44% | 153.1944 | 169.3201
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.52785 | 0.022684 | 0.48248 | 0.57322 | 0.4598 | 0.5959 | 4.297% | 8.595% | 12.89% | 0.50146 | 0.55424
Sb, wt.% | 0.83979 | 0.028525 | 0.78274 | 0.89684 | 0.75421 | 0.92536 | 3.397% | 6.793% | 10.19% | 0.7978 | 0.88178
Sc, ppm | 13.2056 | 0.86954 | 11.4666 | 14.9447 | 10.597 | 15.8142 | 6.585% | 13.17% | 19.75% | 12.5454 | 13.8659
Se, ppm | 2.6423 | 0.8024 | 1.0375 | 4.2471 | 0.23506 | 5.0495 | 30.37% | 60.74% | 91.1% | 2.5102 | 2.7744
Sm, ppm | 6.4231 | 0.44257 | 5.538 | 7.3083 | 5.0954 | 7.7509 | 6.89% | 13.78% | 20.67% | 6.102 | 6.7443
Sn, ppm | 3.5648 | 0.32411 | 2.9166 | 4.213 | 2.5925 | 4.5372 | 9.092% | 18.18% | 27.28% | 3.3866 | 3.7431
Sr, ppm | 99.1289 | 4.597 | 89.9349 | 108.3229 | 85.3379 | 112.9199 | 4.637% | 9.275% | 13.91% | 94.1725 | 104.0854
Ta, ppm | 0.9575 | 0.19752 | 0.56245 | 1.3525 | 0.36492 | 1.5501 | 20.63% | 41.26% | 61.89% | 0.90962 | 1.0054
Tb, ppm | 0.55343 | 0.084973 | 0.38348 | 0.72338 | 0.29851 | 0.80835 | 15.35% | 30.71% | 46.06% | 0.52576 | 0.5811
Te, ppm | 0.071667 | 0.016568 | 0.03853 | 0.1048 | 0.021961 | 0.12137 | 23.12% | 46.24% | 69.36% | 0.068083 | 0.07525
Th, ppm | 15.1315 | 1.2107 | 12.7102 | 17.5529 | 11.4995 | 18.7636 | 8.001% | 16% | 24% | 14.375 | 15.8881
Ti, wt.% | 0.35905 | 0.030555 | 0.29794 | 0.42016 | 0.26738 | 0.45071 | 8.51% | 17.02% | 25.53% | 0.34109 | 0.377
Tl, ppm | 0.85951 | 0.053286 | 0.75294 | 0.96609 | 0.69966 | 1.0194 | 6.2% | 12.4% | 18.6% | 0.81654 | 0.90249
Tm, ppm | 0.223 | 0.030643 | 0.16171 | 0.28429 | 0.13107 | 0.31493 | 13.74% | 27.48% | 41.22% | 0.21185 | 0.23415
U, ppm | 2.785 | 0.19149 | 2.4021 | 3.168 | 2.2106 | 3.3595 | 6.876% | 13.75% | 20.63% | 2.6458 | 2.9243
V, ppm | 91.982 | 5.2713 | 81.4395 | 102.5246 | 76.1682 | 107.7958 | 5.731% | 11.46% | 17.19% | 87.3829 | 96.5811
W, ppm | 2.8289 | 0.34059 | 2.1477 | 3.51 | 1.8071 | 3.8506 | 12.04% | 24.08% | 36.12% | 2.6874 | 2.9703
Y, ppm | 15.1844 | 2.0188 | 11.1468 | 19.222 | 9.128 | 21.2408 | 13.3% | 26.59% | 39.89% | 14.4252 | 15.9437
Yb, ppm | 1.5844 | 0.17743 | 1.2296 | 1.9393 | 1.0521 | 2.1167 | 11.2% | 22.4% | 33.59% | 1.5052 | 1.6636
Zn, ppm | 97.5423 | 5.4431 | 86.6561 | 108.4285 | 81.2129 | 113.8716 | 5.58% | 11.16% | 16.74% | 92.6652 | 102.4194
Zr, ppm | 133.3184 | 8.7577 | 115.8031 | 150.8338 | 107.0454 | 159.5914 | 6.569% | 13.14% | 19.71% | 126.6525 | 139.9843
Aqua Regia Digestion
Ag, ppm | 0.19233 | 0.016586 | 0.15916 | 0.2255 | 0.14257 | 0.24209 | 8.623% | 17.25% | 25.87% | 0.18271 | 0.20195
Al, wt.% | 2.5919 | 0.27899 | 2.0339 | 3.1499 | 1.7549 | 3.4289 | 10.76% | 21.53% | 32.29% | 2.4623 | 2.7215
As, ppm | 306.7984 | 15.7821 | 275.2342 | 338.3627 | 259.452 | 354.1449 | 5.144% | 10.29% | 15.43% | 291.4585 | 322.1384
Ba, ppm | 116.6471 | 6.4243 | 103.7985 | 129.4958 | 97.3742 | 135.9201 | 5.507% | 11.01% | 16.52% | 110.8148 | 122.4795
Be, ppm | 1.2586 | 0.17995 | 0.89871 | 1.6185 | 0.71876 | 1.7984 | 14.3% | 28.59% | 42.89% | 1.1957 | 1.3215
Bi, ppm | 0.36432 | 0.023381 | 0.31756 | 0.41109 | 0.29418 | 0.43447 | 6.418% | 12.84% | 19.25% | 0.34611 | 0.38254
Ca, wt.% | 0.31211 | 0.017315 | 0.27748 | 0.34674 | 0.26017 | 0.36406 | 5.548% | 11.1% | 16.64% | 0.29651 | 0.32772
Cd, ppm | 0.050191 | 0.012271 | 0.025649 | 0.074732 | 0.013378 | 0.087003 | 24.45% | 48.9% | 73.34% | 0.047681 | 0.0527
Ce, ppm | 50.1463 | 10.3756 | 29.3951 | 70.8975 | 19.0195 | 81.2731 | 20.69% | 41.38% | 62.07% | 47.639 | 52.6536
Co, ppm | 13.4206 | 0.70706 | 12.0064 | 14.8347 | 11.2994 | 15.5417 | 5.269% | 10.54% | 15.81% | 12.7495 | 14.0916
Cr, ppm | 98.3039 | 6.1115 | 86.0809 | 110.5269 | 79.9695 | 116.6383 | 6.217% | 12.43% | 18.65% | 93.3887 | 103.2191
Cs, ppm | 7.2424 | 0.2775 | 6.6874 | 7.7974 | 6.4099 | 8.0749 | 3.832% | 7.663% | 11.49% | 6.8803 | 7.6045
Cu, ppm | 31.8903 | 2.216 | 27.4583 | 36.3223 | 25.2423 | 38.5383 | 6.949% | 13.9% | 20.85% | 30.2958 | 33.4848
Fe, wt.% | 3.3821 | 0.19522 | 2.9917 | 3.7725 | 2.7965 | 3.9678 | 5.772% | 11.54% | 17.32% | 3.213 | 3.5512
Ga, ppm | 8.0593 | 0.55144 | 6.9564 | 9.1621 | 6.4049 | 9.7136 | 6.842% | 13.68% | 20.53% | 7.6563 | 8.4622
Ge, ppm | 0.096667 | 0.0066089 | 0.083449 | 0.10988 | 0.07684 | 0.11649 | 6.837% | 13.67% | 20.51% | 0.091833 | 0.1015
Hf, ppm | 0.51173 | 0.053847 | 0.40404 | 0.61942 | 0.35019 | 0.67327 | 10.52% | 21.05% | 31.57% | 0.48614 | 0.53732
Hg, ppm | 0.030738 | 0.0062315 | 0.018275 | 0.043201 | 0.012043 | 0.049433 | 20.27% | 40.55% | 60.82% | 0.029201 | 0.032275
In, ppm | 0.03423 | 0.0047051 | 0.024819 | 0.04364 | 0.020114 | 0.048345 | 13.75% | 27.49% | 41.24% | 0.032518 | 0.035941
K, wt.% | 0.85765 | 0.033689 | 0.79027 | 0.92503 | 0.75658 | 0.95871 | 3.928% | 7.856% | 11.78% | 0.81477 | 0.90053
La, ppm | 23.7008 | 4.0735 | 15.5539 | 31.8477 | 11.4804 | 35.9212 | 17.19% | 34.37% | 51.56% | 22.5157 | 24.8858
Li, ppm | 37.5991 | 4.7131 | 28.1728 | 47.0253 | 23.4597 | 51.7385 | 12.54% | 25.07% | 37.61% | 35.7191 | 39.479
Mg, wt.% | 1.145 | 0.052129 | 1.0407 | 1.2492 | 0.98858 | 1.3014 | 4.553% | 9.106% | 13.66% | 1.0877 | 1.2022
Mn, wt.% | 0.047492 | 0.0020792 | 0.043334 | 0.05165 | 0.041254 | 0.053729 | 4.378% | 8.756% | 13.13% | 0.045117 | 0.049866
Mo, ppm | 1.0946 | 0.10913 | 0.87636 | 1.3129 | 0.76723 | 1.422 | 9.97% | 19.94% | 29.91% | 1.0399 | 1.1494
Na, wt.% | 0.046558 | 0.0044231 | 0.037712 | 0.055404 | 0.033289 | 0.059827 | 9.5% | 19% | 28.5% | 0.04423 | 0.048886
Nb, ppm | 0.53966 | 0.12398 | 0.29169 | 0.78762 | 0.16771 | 0.9116 | 22.97% | 45.95% | 68.92% | 0.51267 | 0.56664
Ni, ppm | 67.7149 | 5.5713 | 56.5722 | 78.8575 | 51.0009 | 84.4288 | 8.228% | 16.46% | 24.68% | 64.3291 | 71.1006
P, wt.% | 0.050105 | 0.0020224 | 0.046061 | 0.05415 | 0.044038 | 0.056173 | 4.036% | 8.073% | 12.11% | 0.0476 | 0.052611
Pb, ppm | 13.5606 | 0.91007 | 11.7405 | 15.3808 | 10.8304 | 16.2909 | 6.711% | 13.42% | 20.13% | 12.8826 | 14.2387
Rb, ppm | 80.8171 | 4.7895 | 71.238 | 90.3962 | 66.4485 | 95.1857 | 5.926% | 11.85% | 17.78% | 76.7763 | 84.858
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.53261 | 0.017185 | 0.49824 | 0.56698 | 0.48105 | 0.58416 | 3.227% | 6.453% | 9.68% | 0.50598 | 0.55924
Sb, wt.% | 0.74463 | 0.082035 | 0.58056 | 0.9087 | 0.49852 | 0.99073 | 11.02% | 22.03% | 33.05% | 0.7074 | 0.78186
Sc, ppm | 6.7834 | 0.63746 | 5.5084 | 8.0583 | 4.871 | 8.6957 | 9.397% | 18.79% | 28.19% | 6.4442 | 7.1225
Se, ppm | 2.0852 | 0.39152 | 1.3022 | 2.8683 | 0.91068 | 3.2598 | 18.78% | 37.55% | 56.33% | 1.981 | 2.1895
Sn, ppm | 1.5634 | 0.10862 | 1.3462 | 1.7807 | 1.2376 | 1.8893 | 6.948% | 13.9% | 20.84% | 1.4853 | 1.6416
Sr, ppm | 32.6002 | 2.0027 | 28.5947 | 36.6056 | 26.5919 | 38.6084 | 6.143% | 12.29% | 18.43% | 30.9702 | 34.2302
Ta, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Te, ppm | 0.053157 | 0.0081637 | 0.03683 | 0.069484 | 0.028666 | 0.077648 | 15.36% | 30.72% | 46.07% | 0.050499 | 0.055815
Th, ppm | 11.3145 | 1.1806 | 8.9533 | 13.6757 | 7.7727 | 14.8563 | 10.43% | 20.87% | 31.3% | 10.7488 | 11.8802
Ti, wt.% | 0.11925 | 0.0075041 | 0.10424 | 0.13426 | 0.096739 | 0.14176 | 6.293% | 12.59% | 18.88% | 0.11329 | 0.12521
Tl, ppm | 0.51524 | 0.030311 | 0.45462 | 0.57586 | 0.42431 | 0.60618 | 5.883% | 11.77% | 17.65% | 0.48948 | 0.541
U, ppm | 1.3848 | 0.17188 | 1.041 | 1.7285 | 0.86914 | 1.9004 | 12.41% | 24.82% | 37.24% | 1.3155 | 1.454
V, ppm | 60.7722 | 2.7215 | 55.3292 | 66.2151 | 52.6078 | 68.9365 | 4.478% | 8.956% | 13.43% | 57.7335 | 63.8108
W, ppm | 0.54297 | 0.077745 | 0.38748 | 0.69846 | 0.30974 | 0.77621 | 14.32% | 28.64% | 42.96% | 0.51582 | 0.57012
Y, ppm | 8.2067 | 0.51834 | 7.17 | 9.2433 | 6.6516 | 9.7617 | 6.316% | 12.63% | 18.95% | 7.7963 | 8.617
Zn, ppm | 86.2117 | 2.8719 | 80.4679 | 91.9556 | 77.5959 | 94.8275 | 3.331% | 6.662% | 9.994% | 81.9011 | 90.5223
Zr, ppm | 19.8408 | 3.0194 | 13.8021 | 25.8796 | 10.7828 | 28.8989 | 15.22% | 30.44% | 45.65% | 18.8488 | 20.8329
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 290
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 2.1156 | 2.0881 | 2.1431 | 2.1072 | 2.124
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.9323 | 1.8722 | 1.9924 | 1.9239 | 1.9407
Peroxide Fusion ICP
Sb, Antimony (wt.%) | 0.79545 | 0.76544 | 0.82547 | 0.77479 | 0.81612
4-Acid Digestion
Ag, Silver (ppm) | 0.21349 | 0.18204 | 0.24494 | 0.19363 | 0.23335
Al, Aluminium (wt.%) | 7.2094 | 6.8747 | 7.5441 | 7.0009 | 7.4178
As, Arsenic (ppm) | 313.437 | 294.3241 | 332.5499 | 300.872 | 326.002
Ba, Barium (ppm) | 738.9002 | 707.6732 | 770.1272 | 721.8212 | 755.9792
Be, Beryllium (ppm) | 2.5085 | 2.3139 | 2.7031 | 2.3867 | 2.6303
Bi, Bismuth (ppm) | 0.36738 | 0.33096 | 0.40381 | 0.34 | 0.39476
Ca, Calcium (wt.%) | 0.52455 | 0.49213 | 0.55696 | 0.50956 | 0.53953
Cd, Cadmium (ppm) | 0.053506 | 0.033822 | 0.073191 | IND | IND
Ce, Cerium (ppm) | 78.7458 | 73.9159 | 83.5756 | 75.1622 | 82.3293
Co, Cobalt (ppm) | 14.1993 | 13.331 | 15.0677 | 13.6082 | 14.7905
Cr, Chromium (ppm) | 122.7493 | 115.2201 | 130.2785 | 117.7885 | 127.7102
Cs, Caesium (ppm) | 10.33 | 9.8089 | 10.851 | 9.9442 | 10.7158
Cu, Copper (ppm) | 32.5699 | 31.1531 | 33.9867 | 31.4566 | 33.6832
Dy, Dysprosium (ppm) | 3.0127 | 2.5112 | 3.5141 | 2.818 | 3.2074
Er, Erbium (ppm) | 1.6002 | 1.3313 | 1.8691 | IND | IND
Eu, Europium (ppm) | 1.141 | 1.0014 | 1.2806 | IND | IND
Fe, Iron (wt.%) | 3.7892 | 3.6221 | 3.9563 | 3.6964 | 3.882
Ga, Gallium (ppm) | 18.6282 | 17.6062 | 19.6501 | 17.9803 | 19.276
Gd, Gadolinium (ppm) | 4.4747 | 3.8774 | 5.0719 | 4.1255 | 4.8238
Ge, Germanium (ppm) | 0.2172 | 0.13546 | 0.29894 | 0.1909 | 0.2435
Hf, Hafnium (ppm) | 3.9792 | 3.7857 | 4.1728 | 3.8364 | 4.1221
Ho, Holmium (ppm) | 0.56167 | 0.44329 | 0.68004 | IND | IND
In, Indium (ppm) | 0.064435 | 0.05465 | 0.074219 | 0.059872 | 0.068998
K, Potassium (wt.%) | 2.8003 | 2.6794 | 2.9212 | 2.7052 | 2.8954
La, Lanthanum (ppm) | 39.5905 | 37.4417 | 41.7394 | 38.3829 | 40.7981
Li, Lithium (ppm) | 51.3773 | 49.3241 | 53.4305 | 49.7969 | 52.9578
Lu, Lutetium (ppm) | 0.23619 | 0.18751 | 0.28487 | 0.20603 | 0.26635
Mg, Magnesium (wt.%) | 1.388 | 1.3274 | 1.4487 | 1.3503 | 1.4258
Mn, Manganese (wt.%) | 0.056358 | 0.053778 | 0.058939 | 0.055337 | 0.05738
Mo, Molybdenum (ppm) | 1.1356 | 0.9701 | 1.301 | 1.0844 | 1.1868
Na, Sodium (wt.%) | 0.61344 | 0.58337 | 0.64351 | 0.5969 | 0.62998
Nb, Niobium (ppm) | 10.1868 | 8.2565 | 12.1171 | 9.4968 | 10.8768
Nd, Neodymium (ppm) | 34.2123 | 31.5593 | 36.8654 | 32.5391 | 35.8855
Ni, Nickel (ppm) | 72.591 | 69.3299 | 75.8522 | 70.4758 | 74.7063
P, Phosphorus (wt.%) | 0.05857 | 0.055365 | 0.061775 | 0.057133 | 0.060007
Pb, Lead (ppm) | 22.8867 | 21.6156 | 24.1577 | 22.0014 | 23.7719
Pr, Praseodymium (ppm) | 9.1335 | 8.4323 | 9.8348 | 8.747 | 9.5201
Rb, Rubidium (ppm) | 161.2573 | 152.7368 | 169.7777 | 155.9812 | 166.5333
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.52785 | 0.51124 | 0.54446 | 0.51285 | 0.54285
Sb, Antimony (wt.%) | 0.83979 | 0.81051 | 0.86906 | 0.82025 | 0.85932
Sc, Scandium (ppm) | 13.2056 | 12.2758 | 14.1355 | 12.7039 | 13.7074
Se, Selenium (ppm) | 2.6423 | 1.9642 | 3.3204 | 2.0556 | 3.229
Sm, Samarium (ppm) | 6.4231 | 5.7904 | 7.0559 | 6.1519 | 6.6944
Sn, Tin (ppm) | 3.5648 | 3.3078 | 3.8219 | 3.4057 | 3.724
Sr, Strontium (ppm) | 99.1289 | 94.2027 | 104.0551 | 96.2118 | 102.046
Ta, Tantalum (ppm) | 0.9575 | 0.81302 | 1.102 | 0.91989 | 0.99511
Tb, Terbium (ppm) | 0.55343 | 0.47499 | 0.63187 | 0.50561 | 0.60124
Te, Tellurium (ppm) | 0.071667 | 0.048953 | 0.09438 | IND | IND
Th, Thorium (ppm) | 15.1315 | 13.9674 | 16.2957 | 14.4097 | 15.8534
Ti, Titanium (wt.%) | 0.35905 | 0.33873 | 0.37936 | 0.34441 | 0.37368
Tl, Thallium (ppm) | 0.85951 | 0.80941 | 0.90962 | 0.82491 | 0.89412
Tm, Thulium (ppm) | 0.223 | 0.17771 | 0.26829 | IND | IND
U, Uranium (ppm) | 2.785 | 2.5827 | 2.9874 | 2.6604 | 2.9097
V, Vanadium (ppm) | 91.982 | 86.8559 | 97.1082 | 88.4287 | 95.5353
W, Tungsten (ppm) | 2.8289 | 2.5099 | 3.1478 | 2.6081 | 3.0496
Y, Yttrium (ppm) | 15.1844 | 13.9297 | 16.4392 | 14.5764 | 15.7925
Yb, Ytterbium (ppm) | 1.5844 | 1.3698 | 1.799 | IND | IND
Zn, Zinc (ppm) | 97.5423 | 93.7153 | 101.3692 | 94.6071 | 100.4775
Zr, Zirconium (ppm) | 133.3184 | 125.7165 | 140.9203 | 129.2078 | 137.429
Aqua Regia Digestion
Ag, Silver (ppm) | 0.19233 | 0.17391 | 0.21076 | 0.17475 | 0.20991
Al, Aluminium (wt.%) | 2.5919 | 2.4388 | 2.745 | 2.5112 | 2.6726
As, Arsenic (ppm) | 306.7984 | 292.562 | 321.0348 | 296.7282 | 316.8686
Ba, Barium (ppm) | 116.6471 | 111.1874 | 122.1069 | 112.2567 | 121.0376
Be, Beryllium (ppm) | 1.2586 | 1.1159 | 1.4013 | 1.199 | 1.3182
Bi, Bismuth (ppm) | 0.36432 | 0.33602 | 0.39263 | 0.34142 | 0.38723
Ca, Calcium (wt.%) | 0.31211 | 0.30126 | 0.32297 | 0.3017 | 0.32252
Cd, Cadmium (ppm) | 0.050191 | 0.037206 | 0.063175 | IND | IND
Ce, Cerium (ppm) | 50.1463 | 42.4943 | 57.7983 | 47.9845 | 52.3082
Co, Cobalt (ppm) | 13.4206 | 12.8614 | 13.9797 | 12.9693 | 13.8718
Cr, Chromium (ppm) | 98.3039 | 94.5001 | 102.1077 | 95.6032 | 101.0046
Cs, Caesium (ppm) | 7.2424 | 6.8937 | 7.5911 | 6.9586 | 7.5262
Cu, Copper (ppm) | 31.8903 | 29.9353 | 33.8453 | 30.581 | 33.1995
Fe, Iron (wt.%) | 3.3821 | 3.2723 | 3.4919 | 3.3006 | 3.4637
Ga, Gallium (ppm) | 8.0593 | 7.5464 | 8.5721 | 7.762 | 8.3565
Ge, Germanium (ppm) | 0.096667 | 0.079566 | 0.11377 | IND | IND
Hf, Hafnium (ppm) | 0.51173 | 0.45013 | 0.57333 | 0.4767 | 0.54676
Hg, Mercury (ppm) | 0.030738 | 0.018784 | 0.042692 | IND | IND
In, Indium (ppm) | 0.03423 | 0.026696 | 0.041764 | 0.025861 | 0.042599
K, Potassium (wt.%) | 0.85765 | 0.83001 | 0.88529 | 0.83415 | 0.88114
La, Lanthanum (ppm) | 23.7008 | 21.1058 | 26.2958 | 22.6691 | 24.7325
Li, Lithium (ppm) | 37.5991 | 34.4816 | 40.7165 | 35.7321 | 39.466
Mg, Magnesium (wt.%) | 1.145 | 1.1122 | 1.1778 | 1.1136 | 1.1763
Mn, Manganese (wt.%) | 0.047492 | 0.046284 | 0.0487 | 0.046414 | 0.04857
Mo, Molybdenum (ppm) | 1.0946 | 0.96398 | 1.2253 | 0.97011 | 1.2192
Na, Sodium (wt.%) | 0.046558 | 0.042591 | 0.050525 | 0.044787 | 0.048329
Nb, Niobium (ppm) | 0.53966 | 0.41417 | 0.66515 | 0.48132 | 0.598
Ni, Nickel (ppm) | 67.7149 | 64.7931 | 70.6366 | 65.9855 | 69.4442
P, Phosphorus (wt.%) | 0.050105 | 0.048999 | 0.051212 | 0.048569 | 0.051642
Pb, Lead (ppm) | 13.5606 | 12.7785 | 14.3428 | 13.038 | 14.0833
Rb, Rubidium (ppm) | 80.8171 | 76.7279 | 84.9063 | 77.8119 | 83.8223
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.53261 | 0.51721 | 0.54801 | 0.51562 | 0.5496
Sb, Antimony (wt.%) | 0.74463 | 0.69814 | 0.79111 | 0.72501 | 0.76424
Sc, Scandium (ppm) | 6.7834 | 6.3207 | 7.246 | 6.5446 | 7.0221
Se, Selenium (ppm) | 2.0852 | 1.6625 | 2.5079 | 1.8483 | 2.3221
Sn, Tin (ppm) | 1.5634 | 1.4457 | 1.6812 | IND | IND
Sr, Strontium (ppm) | 32.6002 | 31.2938 | 33.9065 | 31.3548 | 33.8455
Ta, Tantalum (ppm) | < 0.05 | IND | IND | IND | IND
Te, Tellurium (ppm) | 0.053157 | 0.040643 | 0.065671 | IND | IND
Th, Thorium (ppm) | 11.3145 | 10.3822 | 12.2468 | 10.9055 | 11.7235
Ti, Titanium (wt.%) | 0.11925 | 0.11203 | 0.12647 | 0.11544 | 0.12306
Tl, Thallium (ppm) | 0.51524 | 0.48462 | 0.54586 | 0.4849 | 0.54558
U, Uranium (ppm) | 1.3848 | 1.2666 | 1.503 | 1.3059 | 1.4637
V, Vanadium (ppm) | 60.7722 | 58.6225 | 62.9218 | 58.869 | 62.6753
W, Tungsten (ppm) | 0.54297 | 0.47824 | 0.6077 | 0.49679 | 0.58915
Y, Yttrium (ppm) | 8.2067 | 7.7542 | 8.6591 | 7.9347 | 8.4786
Zn, Zinc (ppm) | 86.2117 | 84.14 | 88.2834 | 83.683 | 88.7404
Zr, Zirconium (ppm) | 19.8408 | 17.3893 | 22.2924 | 18.6197 | 21.062
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 290
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 5 | Pt | ppb | < 5
Peroxide Fusion ICP
Al | wt.% | 7.4126 | Ga | ppm | 14.4084 | S | wt.% | 0.51347
As | ppm | 359.0448 | K | wt.% | 2.8545 | Sc | ppm | 15.3157
Ba | ppm | 742.2599 | La | ppm | 39.8249 | Si | wt.% | 30.6008
Be | ppm | < 5 | Li | ppm | < 5 | Sn | ppm | < 10
Bi | ppm | < 5 | Mg | wt.% | 1.4379 | Sr | ppm | 87.3979
Ca | wt.% | 0.5716 | Mn | wt.% | 0.062091 | Ti | wt.% | 0.43688
Cd | ppm | < 0.5 | Mo | ppm | 2.8313 | V | ppm | 185.9817
Co | ppm | 20.3121 | Nb | ppm | 5.9752 | W | ppm | 72.9748
Cr | ppm | 152.0263 | Ni | ppm | 78.2838 | Y | ppm | 30.7396
Cu | ppm | 61.9988 | P | wt.% | 0.060817 | Zn | ppm | 114.133
Fe | wt.% | 3.9631 | Pb | ppm | 22.5066
4-Acid Digestion
Hg | ppm | 0.10667
Aqua Regia Digestion
B | ppm | 17.324 | Ho | ppm | 0.32175 | Pt | ppb | 2
Dy | ppm | 1.8704 | Lu | ppm | 0.10592 | Sm | ppm | 3.8509
Er | ppm | 0.87017 | Nd | ppm | 19.7263 | Tb | ppm | 0.37296
Eu | ppm | 0.52942 | Pd | ppb | < 10 | Tm | ppm | 0.10983
Gd | ppm | 2.9587 | Pr | ppm | 5.3024 | Yb | ppm | 0.77363
Borate Fusion XRF
Al2O3 | wt.% | 14.0905 | MgO | wt.% | 2.355 | SiO2 | wt.% | 67.935
CaO | wt.% | 0.7415 | MnO | wt.% | 0.0755 | SO3 | wt.% | 1.3065
Fe2O3 | wt.% | 5.575 | P2O5 | wt.% | 0.1325 | TiO2 | wt.% | 0.73
K2O | wt.% | 3.4435 | Sb | wt.% | 0.925
Thermogravimetry
LOI1000 | wt.% | 2.92
Infrared Combustion
C | wt.% | 0.215 | S | wt.% | 0.51
Laser Ablation ICP-MS
Ag | ppm | 0.5 | Hf | ppm | 7.04 | Sm | ppm | 7.59
As | ppm | 360.5 | Ho | ppm | 1.25 | Sn | ppm | 4.4
Ba | ppm | 759.5 | In | ppm | 0.05 | Sr | ppm | 100.8
Be | ppm | 2.3 | La | ppm | 42 | Ta | ppm | 1.175
Bi | ppm | 0.43 | Lu | ppm | 0.45 | Tb | ppm | 1.005
Cd | ppm | < 0.1 | Mn | wt.% | 0.0587 | Te | ppm | 0.15
Ce | ppm | 81.95 | Mo | ppm | 0.9 | Th | ppm | 16.2
Co | ppm | 15.15 | Nb | ppm | 15.2 | Ti | wt.% | 0.4485
Cr | ppm | 134 | Nd | ppm | 37.1 | Tl | ppm | 1.1
Cs | ppm | 11.15 | Ni | ppm | 80 | Tm | ppm | 0.55
Cu | ppm | 40 | Pb | ppm | 30.5 | U | ppm | 3.36
Dy | ppm | 5.945 | Pr | ppm | 9.71 | V | ppm | 103
Er | ppm | 3.585 | Rb | ppm | 171.5 | W | ppm | 3.75
Eu | ppm | 1.39 | Re | ppm | 0.04 | Y | ppm | 31.6
Ga | ppm | 19.4 | Sb | wt.% | 0.939 | Yb | ppm | 3.335
Gd | ppm | 6.07 | Sc | ppm | 14.4 | Zn | ppm | 92.5
Ge | ppm | 1.475 | Se | ppm | < 5 | Zr | ppm | 250.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 290
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
HBr*AAS | hydrobromic acid digestion with atomic absorption spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PPP*XRF | pressed powder pellet with X-ray fluorescence spectroscopy
Text Values:
N.A. | Not Applicable (Lab 2.09)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 290
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate (BV), Manila, Philippines
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Labwest Minerals Analysis, Perth, WA, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 290 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
1 | 1 | ANSLu | 1 | 9 | 9 | 217651 | 0.08222 | 2.2717 | 2.2054
1 | 1 | ANSLu | 1 | 17 | 17 | 217652 | 0.083151 | 2.1304 | 2.1979
1 | 1 | ANSLu | 1 | 16 | 16 | 217653 | 0.085359 | 2.2964 | 2.2068
1 | 1 | ANSLu | 1 | 4 | 4 | 217654 | 0.08489 | 2.2651 | 2.2051
1 | 1 | ANSLu | 1 | 19 | 19 | 217655 | 0.085358 | 2.1453 | 2.1987
1 | 1 | ANSLu | 1 | 1 | 1 | 217656 | 0.086075 | 2.2169 | 2.2025
1 | 1 | ANSLu | 1 | 13 | 13 | 217657 | 0.089486 | 2.1714 | 2.2001
1 | 1 | ANSLu | 1 | 12 | 12 | 217658 | 0.085062 | 2.2085 | 2.202
1 | 1 | ANSLu | 1 | 11 | 11 | 217659 | 0.087532 | 2.2372 | 2.2036
1 | 1 | ANSLu | 1 | 10 | 10 | 217660 | 0.086314 | 2.243 | 2.2039
1 | 1 | ANSLu | 1 | 3 | 3 | 217661 | 0.086417 | 2.1956 | 2.2014
1 | 1 | ANSLu | 1 | 14 | 14 | 217662 | 0.08948 | 2.1965 | 2.2014
1 | 1 | ANSLu | 1 | 20 | 20 | 217663 | 0.089618 | 2.1252 | 2.1976
1 | 1 | ANSLu | 1 | 5 | 5 | 217664 | 0.08043 | 2.1722 | 2.2001
1 | 1 | ANSLu | 1 | 7 | 7 | 217665 | 0.089113 | 2.1739 | 2.2002
1 | 1 | ANSLu | 1 | 2 | 2 | 217666 | 0.082877 | 2.2208 | 2.2027
1 | 1 | ANSLu | 1 | 8 | 8 | 217667 | 0.086849 | 2.1318 | 2.1979
1 | 1 | ANSLu | 1 | 15 | 15 | 217668 | 0.086377 | 2.2768 | 2.2057
1 | 1 | ANSLu | 1 | 18 | 18 | 217669 | 0.088665 | 2.1912 | 2.2011
1 | 1 | ANSLu | 1 | 6 | 6 | 217670 | 0.085126 | 2.1636 | 2.1996
Mean | [6 empty] | 0.08602 | 2.2017 | 2.2017
Median | [7 empty] | 2.1961 | 2.2014
SD | [7 empty] | 0.051464 | 0.0027558
RSD | [7 empty] | 2.337% | 0.1252%
RSD Upscale Factor | [7 empty] | 0.053547
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 290 (Certified Value 2.12 ppm) | [21 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 50g | 25g | 30g | 25g | 40g | 3
| 1 | 1 | 2.2169 | 2.0985 | 2.04 | 2.112 | 2.12 | 2.16 | 2.24 | 2.12 | 2.1 | 2.121 | 2.13 | 2.153 | 2.18 | 2.1713 | 2.062 | 2.15 | 2.037 | 2.06 | 2.07 | 2.14 | 1
| 1 | 2 | 2.2208 | 2.1955 | 1.96 | 2.13 | 2.115 | 2.14 | 2.19 | 2.15 | 2.07 | N.A. | 2.11 | 2.221 | 2.14 | 2.1747 | 2.162 | 2.11 | 2.021 | 1.99 | 2 | 2.16 | #N/A
| 1 | 3 | 2.1956 | 2.066 | 2.01 | 2.072 | 2.119 | 2.14 | 2.19 | 2.17 | 2.15 | 2.137 | 2.06 | 2.198 | 2.13 | 2.1583 | 2.093 | 1.95 | 2.03 | 2 | 1.99 | 2.17 | 16
| 1 | 4 | 2.2651 | 2.147 | 1.94 | 1.845 | 2.117 | 2.17 | 2.2 | 2.18 | 2.11 | 2.133 | 2.08 | 2.192 | 2.12 | 2.1682 | 2.133 | 2.03 | 2.062 | 2.06 | 2.06 | 2.14 | 2.1156
| 1 | 5 | 2.1722 | 2.0775 | 1.95 | 1.875 | 2.122 | 2.15 | 2.19 | 2.18 | 2.1 | 2.158 | 2.37 | 2.2 | 2.13 | 2.1618 | 2.149 | 2.06 | 2.098 | 2.05 | 2.08 | 2.14 | 7
| 1 | 6 | 2.1636 | 2.14 | 2.1 | 2.071 | 2.073 | 2.17 | 2.19 | 2.18 | 2.14 | 2.117 | 2.22 | 2.149 | 2.17 | 2.1659 | 2.295 | 2.16 | 2.027 | 2.05 | 2.04 | 2.16
|  | 7 | 2.1739
|  | 8 | 2.1318
|  | 9 | 2.2717
|  | 10 | 2.243
|  | 11 | 2.2372
|  | 12 | 2.2085
|  | 13 | 2.1714
|  | 14 | 2.1965
|  | 15 | 2.2768
|  | 16 | 2.2964
|  | 17 | 2.1304
|  | 18 | 2.1912
|  | 19 | 2.1453
|  | 20 | 2.1252
| Mean |  | 2.2017 | 2.1207 | 2 | 2.0175 | 2.111 | 2.155 | 2.2 | 2.1633 | 2.1117 | 2.1332 | 2.1617 | 2.1855 | 2.145 | 2.1667 | 2.149 | 2.0767 | 2.0458 | 2.035 | 2.04 | 2.1517
| Median |  | 2.1961 | 2.1193 | 1.985 | 2.0715 | 2.118 | 2.155 | 2.19 | 2.175 | 2.105 | 2.133 | 2.12 | 2.195 | 2.135 | 2.1671 | 2.141 | 2.085 | 2.0335 | 2.05 | 2.05 | 2.15
| Std Dev. |  | 0.051464 | 0.049011 | 0.06229 | 0.12448 | 0.018772 | 0.013784 | 0.02 | 0.024221 | 0.029269 | 0.016131 | 0.11618 | 0.028487 | 0.02429 | 0.0060229 | 0.080556 | 0.079917 | 0.029267 | 0.031464 | 0.037417 | 0.013292
| Rel.Std.Dev. |  | 2.337% | 2.311% | 3.114% | 6.17% | 0.8893% | 0.6396% | 0.9091% | 1.12% | 1.386% | 0.7562% | 5.374% | 1.303% | 1.132% | 0.278% | 3.749% | 3.848% | 1.431% | 1.546% | 1.834% | 0.6177%
| PDM3 |  | 4.069% | 0.244% | -5.464% | -4.636% | -0.2168% | 1.863% | 3.99% | 2.257% | -0.1853% | 0.8325% | 2.178% | 3.305% | 1.39% | 2.417% | 1.579% | -1.84% | -3.297% | -3.809% | -3.573% | 1.705%
| Z-Score (Absolute) |  | NA | 0.28 | 2.98 | 2.59 | 0.5 | 0.49 | 1.49 | 0.67 | 0.48 | 0 | 0.64 | 1.17 | 0.26 | 0.75 | 0.35 | 1.27 | 1.96 | 2.2 | 2.09 | 0.41
| Analytical results for Pd in OREAS 290 (Indicative Value < 5 ppb) | [21 empty] | Indicative
Pd | Round | Replicate | Lab | [19 empty] | 1
| No | No | 2.17 | [19 empty] | ppb
|  |  | FA*OES | [19 empty] | 2
|  |  | 30g | [19 empty] | 2
| 1 | 1 | < 5 | [19 empty] | 1
| 1 | 2 | < 5 | [19 empty] | 1
| 1 | 3 | < 5 | [19 empty] | 16
| 1 | 4 | < 5 | [19 empty] | < 5
| 1 | 5 | < 5 | [19 empty] | 7
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 290 (Indicative Value < 5 ppb) | [21 empty] | Indicative
Pt | Round | Replicate | Lab | [19 empty] | 1
| No | No | 2.17 | [19 empty] | ppb
|  |  | FA*OES | [19 empty] | 2
|  |  | 30g | [19 empty] | 2
| 1 | 1 | < 5 | [19 empty] | 1
| 1 | 2 | < 5 | [19 empty] | 1
| 1 | 3 | < 5 | [19 empty] | 16
| 1 | 4 | < 5 | [19 empty] | < 5
| 1 | 5 | < 5 | [19 empty] | 7
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 290 (Certified Value 1.93 ppm) | [14 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | ppm
|  |  | INAA | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*OES | 2
|  |  | 0.085g | 15g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 30g | 30g | 25g | 3
| 1 | 1 | 2.2169 | 2.036 | 1.9566 | 1.83 | 1.87 | 1.95 | 1.88 | > 0.5 | 2.07 | 1.86 | 1.58 | 1.9 | 1.94 | 1
| 1 | 2 | 2.2208 | 2.042 | 1.9324 | 1.57 | 2 | 1.96 | 1.86 | > 0.5 | 2.07 | 1.82 | 1.61 | 1.9 | 1.94 | #N/A
| 1 | 3 | 2.1956 | 1.985 | 2.1568 | 2.03 | 1.78 | 1.96 | 1.84 | > 0.5 | 2.05 | 1.89 | 1.52 | 1.9 | 1.97 | 16
| 1 | 4 | 2.2651 | 1.975 | 2.0556 | 2.08 | 1.78 | 1.95 | 1.82 | > 0.5 | 2.03 | 1.84 | 1.65 | 1.9 | 1.96 | 1.9323
| 1 | 5 | 2.1722 | 2.081 | 1.9604 | 1.78 | 1.76 | 2 | 1.9 | > 0.5 | 2.07 | 1.82 | 1.49 | 1.9 | 1.94 | 9
| 1 | 6 | 2.1636 | 2.083 | 1.9365 | 1.96 | 1.76 | 1.99 | 1.86 | > 0.5 | 2.08 | 1.89 | 1.56 | 1.9 | 1.96
|  | 7 | 2.1739
|  | 8 | 2.1318
|  | 9 | 2.2717
|  | 10 | 2.243
|  | 11 | 2.2372
|  | 12 | 2.2085
|  | 13 | 2.1714
|  | 14 | 2.1965
|  | 15 | 2.2768
|  | 16 | 2.2964
|  | 17 | 2.1304
|  | 18 | 2.1912
|  | 19 | 2.1453
|  | 20 | 2.1252
| Mean |  | 2.2017 | 2.0337 | 1.9997 | 1.875 | 1.825 | 1.9683 | 1.86 |  | 2.0617 | 1.8533 | 1.5683 | 1.9 | 1.9517
| Median |  | 2.1961 | 2.039 | 1.9585 | 1.895 | 1.78 | 1.96 | 1.86 |  | 2.07 | 1.85 | 1.57 | 1.9 | 1.95
| Std Dev. |  | 0.051464 | 0.045955 | 0.089141 | 0.18833 | 0.095026 | 0.02137 | 0.028284 |  | 0.018348 | 0.032042 | 0.058452 | 2.4324e-16 | 0.013292
| Rel.Std.Dev. |  | 2.337% | 2.26% | 4.458% | 10.04% | 5.207% | 1.086% | 1.521% |  | 0.89% | 1.729% | 3.727% | 1.28e-14% | 0.681%
| PDM3 |  | 13.94% | 5.246% | 3.489% | -2.965% | -5.553% | 1.865% | -3.741% |  | 6.695% | -4.087% | -18.84% | -1.671% | 1.002%
| Z-Score (Absolute) |  | NA | 1.32 | 0.98 | 0.25 | 0.74 | 0.67 | 0.39 | NA | 1.6 | 0.46 | 3.27 | 0 | 0.51
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Al in OREAS 290 (Indicative Value 7.41 wt.%) | [15 empty] | Indicative
Al | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | 7.72 | 7.2153 | [12 empty] | 1
| 1 | 2 | 7.7 | 7.2153 | [12 empty] | 3
| 1 | 3 | 7.7 | 7.1582 | [12 empty] | 16
| 1 | 4 | 7.61 | 7.0965 | [12 empty] | 7.4126
| 1 | 5 | 7.6 | 7.2152 | [12 empty] | 9
| 1 | 6 | 7.52 | 7.201
| Mean |  | 7.6417 | 7.1836
| Median |  | 7.655 | 7.2081
| Std Dev. |  | 0.078081 | 0.048035
| Rel.Std.Dev. |  | 1.022% | 0.6687%
| PDM3 |  | 3.09% | -3.09%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for As in OREAS 290 (Indicative Value 359 ppm) | [15 empty] | Indicative
As | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 368 | 361.3342 | [12 empty] | 1
| 1 | 2 | 350 | 356.5075 | [12 empty] | 4
| 1 | 3 | 382 | 358.9532 | [12 empty] | 16
| 1 | 4 | 389 | 350.8535 | [12 empty] | 359.0448
| 1 | 5 | 354 | 355.3 | [12 empty] | 10
| 1 | 6 | 326 | 356.5897
| Mean |  | 361.5 | 356.5897
| Median |  | 361 | 356.5486
| Std Dev. |  | 23.0976 | 3.5404
| Rel.Std.Dev. |  | 6.389% | 0.9928%
| PDM3 |  | 0.6838% | -0.6838%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ba in OREAS 290 (Indicative Value 742 ppm) | [15 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 743 | 745.78 | [12 empty] | 1
| 1 | 2 | 743 | 743.5527 | [12 empty] | 5
| 1 | 3 | 751 | 742.4325 | [12 empty] | 16
| 1 | 4 | 744 | 740.59 | [12 empty] | 742.2599
| 1 | 5 | 741 | 740.5647 | [12 empty] | 11
| 1 | 6 | 734 | 738.1992
| Mean |  | 742.6667 | 741.8532
| Median |  | 743 | 741.5113
| Std Dev. |  | 5.465 | 2.6552
| Rel.Std.Dev. |  | 0.7359% | 0.3579%
| PDM3 |  | 0.0548% | -0.0548%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Be in OREAS 290 (Indicative Value < 5 ppm) | [15 empty] | Indicative
Be | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 2
2
| 1 | 1 | < 5 | [13 empty] | 1
| 1 | 2 | < 5 | [13 empty] | 6
| 1 | 3 | < 5 | [13 empty] | 16
| 1 | 4 | < 5 | [13 empty] | < 5
| 1 | 5 | < 5 | [13 empty] | 12
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 290 (Indicative Value < 5 ppm) | [15 empty] | Indicative
Bi | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 2
2
| 1 | 1 | < 5 | [13 empty] | 1
| 1 | 2 | < 5 | [13 empty] | 7
| 1 | 3 | < 5 | [13 empty] | 16
| 1 | 4 | < 5 | [13 empty] | < 5
| 1 | 5 | < 5 | [13 empty] | 13
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ca in OREAS 290 (Indicative Value 0.572 wt.%) | [15 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.67 | 0.462 | [12 empty] | 1
| 1 | 2 | 0.67 | 0.4515 | [12 empty] | 8
| 1 | 3 | 0.71 | 0.483 | [12 empty] | 16
| 1 | 4 | 0.63 | 0.46925 | [12 empty] | 0.5716
| 1 | 5 | 0.67 | 0.47346 | [12 empty] | 14
| 1 | 6 | 0.69 | 0.48
| Mean |  | 0.67333 | 0.46987
| Median |  | 0.67 | 0.47135
| Std Dev. |  | 0.026583 | 0.011727
| Rel.Std.Dev. |  | 3.948% | 2.496%
| PDM3 |  | 17.8% | -17.8%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cd in OREAS 290 (Indicative Value < 0.5 ppm) | [15 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | < 100 | < 0.5 | [12 empty] | 1
| 1 | 2 | < 100 | < 0.5 | [12 empty] | 9
| 1 | 3 | < 100 | < 0.5 | [12 empty] | 16
| 1 | 4 | < 100 | < 0.5 | [12 empty] | < 0.5
| 1 | 5 | < 100 | < 0.5 | [12 empty] | 15
| 1 | 6 | < 100 | < 0.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Co in OREAS 290 (Indicative Value 20.3 ppm) | [15 empty] | Indicative
Co | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 1
1
| 1 | 1 | 19 | 20.305 | [12 empty] | 1
| 1 | 2 | 18 | 20.47 | [12 empty] | 10
| 1 | 3 | 22 | 20.71 | [12 empty] | 16
| 1 | 4 | 22 | 20.5 | [12 empty] | 20.3121
| 1 | 5 | 20 | 20.365 | [12 empty] | 16
| 1 | 6 | 20 | 20.395
| Mean |  | 20.1667 | 20.4575
| Median |  | 20 | 20.4325
| Std Dev. |  | 1.6021 | 0.14243
| Rel.Std.Dev. |  | 7.944% | 0.6962%
| PDM3 |  | -0.7159% | 0.7159%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cr in OREAS 290 (Indicative Value 152 ppm) | [15 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 173 | 131.3869 | [12 empty] | 1
| 1 | 2 | 167 | 136.57 | [12 empty] | 11
| 1 | 3 | 175 | 135.6017 | [12 empty] | 16
| 1 | 4 | 158 | 136.0664 | [12 empty] | 152.0263
| 1 | 5 | 168 | 138.94 | [12 empty] | 17
| 1 | 6 | 170 | 134.75
| Mean |  | 168.5 | 135.5525
| Median |  | 169 | 135.8341
| Std Dev. |  | 5.9582 | 2.4808
| Rel.Std.Dev. |  | 3.536% | 1.83%
| PDM3 |  | 10.84% | -10.84%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 290 (Indicative Value 62 ppm) | [15 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 80 | 36.4275 | [12 empty] | 1
| 1 | 2 | 109 | 36.67 | [12 empty] | 12
| 1 | 3 | 132 | 36.518 | [12 empty] | 16
| 1 | 4 | 74 | 34.9131 | [12 empty] | 61.9988
| 1 | 5 | 69 | 35.8036 | [12 empty] | 18
| 1 | 6 | 64 | 35.6535
| Mean |  | 88 | 35.9976
| Median |  | 77 | 36.1156
| Std Dev. |  | 26.7357 | 0.66933
| Rel.Std.Dev. |  | 30.38% | 1.859%
| PDM3 |  | 41.94% | -41.94%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Fe in OREAS 290 (Indicative Value 3.96 wt.%) | [15 empty] | Indicative
Fe | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | 4.09 | 3.8282 | [12 empty] | 1
| 1 | 2 | 4.29 | 3.8218 | [12 empty] | 13
| 1 | 3 | 4.21 | 3.81 | [12 empty] | 16
| 1 | 4 | 4.08 | 3.819 | [12 empty] | 3.9631
| 1 | 5 | 4.14 | 3.8054 | [12 empty] | 19
| 1 | 6 | 3.86 | 3.8024
| Mean |  | 4.1117 | 3.8145
| Median |  | 4.115 | 3.8145
| Std Dev. |  | 0.14634 | 0.010103
| Rel.Std.Dev. |  | 3.559% | 0.2649%
| PDM3 |  | 3.749% | -3.749%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ga in OREAS 290 (Indicative Value 14.4 ppm) | [15 empty] | Indicative
Ga | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 1
1
| 1 | 1 | 14.535 | [13 empty] | 1
| 1 | 2 | 14.172 | [13 empty] | 46
| 1 | 3 | 14.592 | [13 empty] | 16
| 1 | 4 | 14.205 | [13 empty] | 14.4084
| 1 | 5 | 14.394 | [13 empty] | 9
| 1 | 6 | 14.5523
| Mean |  | 14.4084
| Median |  | 14.4645
| Std Dev. |  | 0.18324
| Rel.Std.Dev. |  | 1.272%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 290 (Indicative Value 2.85 wt.%) | [15 empty] | Indicative
K | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | 2.8 | 2.864 | [12 empty] | 1
| 1 | 2 | 2.9 | 2.853 | [12 empty] | 4
| 1 | 3 | 3 | 2.816 | [12 empty] | 16
| 1 | 4 | 2.8 | 2.86 | [12 empty] | 2.8545
| 1 | 5 | 2.9 | 2.823 | [12 empty] | 10
| 1 | 6 | 2.8 | 2.8378
| Mean |  | 2.8667 | 2.8423
| Median |  | 2.85 | 2.8454
| Std Dev. |  | 0.08165 | 0.019914
| Rel.Std.Dev. |  | 2.848% | 0.7006%
| PDM3 |  | 0.4267% | -0.4267%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for La in OREAS 290 (Indicative Value 39.8 ppm) | [15 empty] | Indicative
La | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 1
1
| 1 | 1 | 39.79 | [13 empty] | 1
| 1 | 2 | 40.95 | [13 empty] | 5
| 1 | 3 | 40.345 | [13 empty] | 16
| 1 | 4 | 39.4192 | [13 empty] | 39.8249
| 1 | 5 | 39.135 | [13 empty] | 11
| 1 | 6 | 39.31
| Mean |  | 39.8249
| Median |  | 39.6046
| Std Dev. |  | 0.69904
| Rel.Std.Dev. |  | 1.755%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 290 (Indicative Value < 5 ppm) | [15 empty] | Indicative
Li | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 2
2
| 1 | 1 | < 5 | [13 empty] | 1
| 1 | 2 | < 5 | [13 empty] | 6
| 1 | 3 | < 5 | [13 empty] | 16
| 1 | 4 | < 5 | [13 empty] | < 5
| 1 | 5 | < 5 | [13 empty] | 12
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Mg in OREAS 290 (Indicative Value 1.44 wt.%) | [15 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | 1.53 | 1.3587 | [12 empty] | 1
| 1 | 2 | 1.5 | 1.3569 | [12 empty] | 7
| 1 | 3 | 1.52 | 1.3464 | [12 empty] | 16
| 1 | 4 | 1.53 | 1.3502 | [12 empty] | 1.4379
| 1 | 5 | 1.51 | 1.3764 | [12 empty] | 13
| 1 | 6 | 1.5 | 1.3764
| Mean |  | 1.515 | 1.3608
| Median |  | 1.515 | 1.3578
| Std Dev. |  | 0.013784 | 0.012862
| Rel.Std.Dev. |  | 0.9098% | 0.9451%
| PDM3 |  | 5.361% | -5.361%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mn in OREAS 290 (Indicative Value 0.062 wt.%) | [15 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.0591 | 0.065322 | [12 empty] | 1
| 1 | 2 | 0.0589 | 0.065156 | [12 empty] | 8
| 1 | 3 | 0.0595 | 0.064912 | [12 empty] | 16
| 1 | 4 | 0.0588 | 0.065975 | [12 empty] | 0.062091
| 1 | 5 | 0.059 | 0.065159 | [12 empty] | 14
| 1 | 6 | 0.0582 | 0.065072
| Mean |  | 0.058917 | 0.065266
| Median |  | 0.05895 | 0.065157
| Std Dev. |  | 0.00042622 | 0.00037213
| Rel.Std.Dev. |  | 0.7234% | 0.5702%
| PDM3 |  | -5.113% | 5.113%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mo in OREAS 290 (Indicative Value 2.83 ppm) | [15 empty] | Indicative
Mo | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 2
2
| 1 | 1 | 2.8868 | [13 empty] | 1
| 1 | 2 | 2.7727 | [13 empty] | 9
| 1 | 3 | 2.8463 | [13 empty] | 16
| 1 | 4 | 2.803 | [13 empty] | 2.8313
| 1 | 5 | 2.8594 | [13 empty] | 15
| 1 | 6 | 2.8195
| Mean |  | 2.8313
| Median |  | 2.8329
| Std Dev. |  | 0.041184
| Rel.Std.Dev. |  | 1.455%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 290 (Indicative Value 5.98 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 2
2
| 1 | 1 | 5.886 | [13 empty] | 1
| 1 | 2 | 6.057 | [13 empty] | 10
| 1 | 3 | 5.7945 | [13 empty] | 16
| 1 | 4 | 5.97 | [13 empty] | 5.9752
| 1 | 5 | 5.6985 | [13 empty] | 16
| 1 | 6 | 6.4455
| Mean |  | 5.9752
| Median |  | 5.928
| Std Dev. |  | 0.26271
| Rel.Std.Dev. |  | 4.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 290 (Indicative Value 78 ppm) | [15 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 87 | 73.2824 | [12 empty] | 1
| 1 | 2 | 87 | 73.9674 | [12 empty] | 11
| 1 | 3 | 83 | 73.37 | [12 empty] | 16
| 1 | 4 | 81 | 74.9308 | [12 empty] | 78.2838
| 1 | 5 | 77 | 72.7249 | [12 empty] | 17
| 1 | 6 | 82 | 74.1306
| Mean |  | 82.8333 | 73.7343
| Median |  | 82.5 | 73.6687
| Std Dev. |  | 3.8166 | 0.77446
| Rel.Std.Dev. |  | 4.608% | 1.05%
| PDM3 |  | 5.812% | -5.812%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for P in OREAS 290 (Indicative Value 0.061 wt.%) | [15 empty] | Indicative
P | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | wt.%
|  |  | PF*OES | [13 empty] | 3
3
| 1 | 1 | 0.060048 | [13 empty] | 1
| 1 | 2 | 0.060376 | [13 empty] | 12
| 1 | 3 | 0.061054 | [13 empty] | 16
| 1 | 4 | 0.060819 | [13 empty] | 0.060817
| 1 | 5 | 0.061026 | [13 empty] | 18
| 1 | 6 | 0.061577
| Mean |  | 0.060817
| Median |  | 0.060923
| Std Dev. |  | 0.0005412
| Rel.Std.Dev. |  | 0.8899%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 290 (Indicative Value 22.5 ppm) | [15 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 1
1
| 1 | 1 | 61 | 20.3585 | [12 empty] | 1
| 1 | 2 | < 50 | 20.5 | [12 empty] | 13
| 1 | 3 | < 50 | 20.4858 | [12 empty] | 16
| 1 | 4 | < 50 | 19.16 | [12 empty] | 22.5066
| 1 | 5 | < 50 | 19.915 | [12 empty] | 19
| 1 | 6 | < 50 | 19.66
| Mean |  | 61 | 20.0132
| Median |  | 61 | 20.1367
| Std Dev. |  |  | 0.53701
| Rel.Std.Dev. |  |  | 2.683%
| PDM3 |  | 171% | -11.08%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 290 (Indicative Value 0.513 wt.%) | [15 empty] | Indicative
S | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | wt.%
|  |  | PF*OES | [13 empty] | 3
3
| 1 | 1 | 0.51 | [13 empty] | 1
| 1 | 2 | 0.515 | [13 empty] | 30
| 1 | 3 | 0.51083 | [13 empty] | 16
| 1 | 4 | 0.515 | [13 empty] | 0.51347
| 1 | 5 | 0.52 | [13 empty] | 9
| 1 | 6 | 0.51
| Mean |  | 0.51347
| Median |  | 0.51292
| Std Dev. |  | 0.0039587
| Rel.Std.Dev. |  | 0.771%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 290 (Certified Value 0.795 wt.%) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PPP*XRF | PF*MS | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.7398 | 0.77 | 0.81 | 0.79 | 0.81 | 0.8077 | 0.82 | 0.87807 | 0.82155 | 0.82 | 0.77699 | 0.84 | 0.77594 | 0.755 | 1
| 1 | 2 | 0.732 | 0.78 | 0.8 | 0.81 | 0.79 | 0.7746 | 0.84 | 0.86932 | 0.79544 | 0.8 | 0.77989 | 0.84 | 0.7761 | 0.774 | 26
| 1 | 3 | 0.7676 | 0.77 | 0.83 | 0.8 | 0.78 | 0.7681 | 0.85 | 0.86504 | 0.81271 | 0.82 | 0.78866 | 0.84 | 0.78643 | 0.783 | 16
| 1 | 4 | 0.7098 | 0.77 | 0.83 | 0.81 | 0.79 | 0.773 | 0.82 | 0.87456 | 0.80539 | 0.81 | 0.78629 | 0.86 | 0.77438 | 0.799 | 0.79545
| 1 | 5 | 0.7271 | 0.77 | 0.79 | 0.81 | 0.79 | 0.781 | 0.85 | 0.88072 | 0.79956 | 0.81 | 0.79513 | 0.85 | 0.76841 | 0.843 | 11
| 1 | 6 | 0.7572 | 0.79 | 0.81 | 0.81 | 0.81 | 0.7566 | 0.8 | 0.86702 | 0.7896 | 0.81 | 0.81004 | 0.85 | 0.75443 | 0.787
| Mean |  | 0.73892 | 0.775 | 0.81167 | 0.805 | 0.795 | 0.77683 | 0.83 | 0.87245 | 0.80404 | 0.81167 | 0.7895 | 0.84667 | 0.77261 | 0.79017
| Median |  | 0.7359 | 0.77 | 0.81 | 0.81 | 0.79 | 0.7738 | 0.83 | 0.87194 | 0.80248 | 0.81 | 0.78747 | 0.845 | 0.77516 | 0.785
| Std Dev. |  | 0.020943 | 0.0083666 | 0.016021 | 0.0083666 | 0.012247 | 0.017178 | 0.02 | 0.0062991 | 0.011714 | 0.0075277 | 0.011945 | 0.008165 | 0.010637 | 0.029762
| Rel.Std.Dev. |  | 2.834% | 1.08% | 1.974% | 1.039% | 1.541% | 2.211% | 2.41% | 0.722% | 1.457% | 0.9274% | 1.513% | 0.9644% | 1.377% | 3.767%
| PDM3 |  | -7.108% | -2.571% | 2.038% | 1.2% | -0.05705% | -2.341% | 4.343% | 9.68% | 1.08% | 2.038% | -0.7485% | 6.438% | -2.871% | -0.6647%
| Z-Score (Absolute) |  | 2.35 | 0.95 | 0.47 | 0.21 | 0.18 | 0.88 | 1.18 | 2.82 | 0.18 | 0.47 | 0.39 | 1.83 | 1.04 | 0.36
| Analytical results for Sc in OREAS 290 (Indicative Value 15.3 ppm) | [15 empty] | Indicative
Sc | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 1
1
| 1 | 1 | 16.04 | [13 empty] | 1
| 1 | 2 | 14.92 | [13 empty] | 47
| 1 | 3 | 14.765 | [13 empty] | 16
| 1 | 4 | 15.055 | [13 empty] | 15.3157
| 1 | 5 | 15.61 | [13 empty] | 10
| 1 | 6 | 15.5042
| Mean |  | 15.3157
| Median |  | 15.2796
| Std Dev. |  | 0.48468
| Rel.Std.Dev. |  | 3.165%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Si in OREAS 290 (Indicative Value 30.6 wt.%) | [15 empty] | Indicative
Si | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | > 15 | 30.4172 | [12 empty] | 1
| 1 | 2 | > 15 | 30.5023 | [12 empty] | 5
| 1 | 3 | > 15 | 30.4895 | [12 empty] | 16
| 1 | 4 | > 15 | 30.5029 | [12 empty] | 30.6008
| 1 | 5 | > 15 | 30.7955 | [12 empty] | 11
| 1 | 6 | > 15 | 30.8975
| Mean |  |  | 30.6008
| Median |  |  | 30.5026
| Std Dev. |  |  | 0.19562
| Rel.Std.Dev. |  |  | 0.6393%
| PDM3 |  |  | 1.11e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Sn in OREAS 290 (Indicative Value < 10 ppm) | [15 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 1
1
| 1 | 1 | 252 | < 10 | [12 empty] | 1
| 1 | 2 | 262 | < 10 | [12 empty] | 6
| 1 | 3 | < 100 | < 10 | [12 empty] | 16
| 1 | 4 | < 100 | < 10 | [12 empty] | < 10
| 1 | 5 | < 100 | < 10 | [12 empty] | 12
| 1 | 6 | < 100 | < 10
| Mean |  | 257
| Median |  | 257
| Std Dev. |  | 7.0711
| Rel.Std.Dev. |  | 2.751%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sr in OREAS 290 (Indicative Value 87 ppm) | [15 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 99 | 75.1253 | [12 empty] | 1
| 1 | 2 | 100 | 74.3695 | [12 empty] | 7
| 1 | 3 | 102 | 76.1757 | [12 empty] | 16
| 1 | 4 | 98 | 75.7125 | [12 empty] | 87.3979
| 1 | 5 | 100 | 74.6779 | [12 empty] | 13
| 1 | 6 | 100 | 73.7133
| Mean |  | 99.8333 | 74.9624
| Median |  | 100 | 74.9016
| Std Dev. |  | 1.3292 | 0.90047
| Rel.Std.Dev. |  | 1.331% | 1.201%
| PDM3 |  | 14.23% | -14.23%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ti in OREAS 290 (Indicative Value 0.437 wt.%) | [15 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.45 | 0.43125 | [12 empty] | 1
| 1 | 2 | 0.45 | 0.425 | [12 empty] | 8
| 1 | 3 | 0.45 | 0.425 | [12 empty] | 16
| 1 | 4 | 0.44 | 0.425 | [12 empty] | 0.43688
| 1 | 5 | 0.45 | 0.43125 | [12 empty] | 14
| 1 | 6 | 0.44 | 0.425
| Mean |  | 0.44667 | 0.42708
| Median |  | 0.45 | 0.425
| Std Dev. |  | 0.005164 | 0.0032275
| Rel.Std.Dev. |  | 1.156% | 0.7557%
| PDM3 |  | 2.241% | -2.241%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V in OREAS 290 (Indicative Value 186 ppm) | [15 empty] | Indicative
V | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 241 | 98.8707 | [12 empty] | 1
| 1 | 2 | 472 | 99.3868 | [12 empty] | 9
| 1 | 3 | 334 | 99.27 | [12 empty] | 16
| 1 | 4 | 237 | 98.5062 | [12 empty] | 185.9817
| 1 | 5 | 303 | 99.447 | [12 empty] | 15
| 1 | 6 | < 100 | 99.2992
| Mean |  | 317.4 | 99.13
| Median |  | 303 | 99.2846
| Std Dev. |  | 95.7669 | 0.36631
| Rel.Std.Dev. |  | 30.17% | 0.3695%
| PDM3 |  | 70.66% | -46.7%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for W in OREAS 290 (Indicative Value 73 ppm) | [15 empty] | Indicative
W | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | < 100 | 73.01 | [12 empty] | 1
| 1 | 2 | < 100 | 71.235 | [12 empty] | 10
| 1 | 3 | < 100 | 74.695 | [12 empty] | 16
| 1 | 4 | < 100 | 72.2863 | [12 empty] | 72.9748
| 1 | 5 | < 100 | 73.69 | [12 empty] | 16
| 1 | 6 | < 100 | 72.9325
| Mean |  |  | 72.9748
| Median |  |  | 72.9712
| Std Dev. |  |  | 1.1811
| Rel.Std.Dev. |  |  | 1.618%
| PDM3 |  |  | -5.551e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Y in OREAS 290 (Indicative Value 30.7 ppm) | [15 empty] | Indicative
Y | Round | Replicate | Lab | [13 empty] | 1
| No | No | 2.19 | [13 empty] | ppm
|  |  | PF*OES | [13 empty] | 1
1
| 1 | 1 | 30.527 | [13 empty] | 1
| 1 | 2 | 30.3842 | [13 empty] | 11
| 1 | 3 | 31.107 | [13 empty] | 16
| 1 | 4 | 30.521 | [13 empty] | 30.7396
| 1 | 5 | 31.6095 | [13 empty] | 17
| 1 | 6 | 30.2888
| Mean |  | 30.7396
| Median |  | 30.524
| Std Dev. |  | 0.51269
| Rel.Std.Dev. |  | 1.668%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 290 (Indicative Value 114 ppm) | [15 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.19 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | [12 empty] | 0
0
| 1 | 1 | 155 | 100.07 | [12 empty] | 1
| 1 | 2 | 127 | 104.2608 | [12 empty] | 12
| 1 | 3 | 121 | 102.61 | [12 empty] | 16
| 1 | 4 | 115 | 107.69 | [12 empty] | 114.133
| 1 | 5 | 114 | 103.095 | [12 empty] | 18
| 1 | 6 | 117 | 102.87
| Mean |  | 124.8333 | 103.4326
| Median |  | 119 | 102.9825
| Std Dev. |  | 15.5231 | 2.499
| Rel.Std.Dev. |  | 12.44% | 2.416%
| PDM3 |  | 9.375% | -9.375%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 290 (Certified Value 0.213 ppm) | [19 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.64 | 0.18 | 0.174 | 0.25 | 0.17 | 0.19 | 0.19 | 0.34 | < 2 | 0.26 | < 2 | 0.23888 | 0.19 | 0.3 | 0.4 | 0.19 | 0.8975 | 0.21 | 1
| 1 | 2 | 0.51 | 0.22 | 0.167 | 0.25 | 0.19 | 0.19 | 0.18 | 0.27 | < 2 | 0.26 | < 2 | 0.25678 | 0.19 | 0.4 | 0.3 | 0.19 | 0.9 | 0.2 | 32
| 1 | 3 | 0.67 | 0.21 | 0.166 | 0.21 | 0.17 | 0.2 | 0.21 | 0.24 | < 2 | 0.25 | < 2 | 0.19902 | 0.21 | 0.3 | 0.4 | 0.18 | 0.872 | 0.17 | 16
| 1 | 4 | 0.59 | 0.19 | 0.18 | 0.22 | 0.19 | 0.21 | 0.2 | 0.22 | < 2 | 0.23 | < 2 | 0.18152 | 0.23 | 0.3 | 0.4 | 0.2 | 0.86375 | 0.18 | 0.21349
| 1 | 5 | 0.69 | 0.21 | 0.171 | 0.22 | 0.15 | 0.18 | 0.21 | 0.22 | < 2 | 0.2 | < 2 | 0.28213 | 0.19 | 0.3 | 0.3 | 0.18 | 0.87225 | 0.19 | 13
| 1 | 6 | 0.69 | 0.21 | 0.175 | 0.22 | 0.18 | 0.2 | 0.18 | 0.55 | < 2 | 0.24 | < 2 | 0.27912 | 0.19 | 0.3 | 0.4 | 0.18 | 0.865 | 0.18
| Mean |  | 0.63167 | 0.20333 | 0.17217 | 0.22833 | 0.175 | 0.195 | 0.195 | 0.30667 |  | 0.24 |  | 0.23958 | 0.2 | 0.31667 | 0.36667 | 0.18667 | 0.87842 | 0.18833
| Median |  | 0.655 | 0.21 | 0.1725 | 0.22 | 0.175 | 0.195 | 0.195 | 0.255 |  | 0.245 |  | 0.24783 | 0.19 | 0.3 | 0.4 | 0.185 | 0.87213 | 0.185
| Std Dev. |  | 0.070545 | 0.015055 | 0.0052694 | 0.017224 | 0.015166 | 0.010488 | 0.013784 | 0.12738 |  | 0.022804 |  | 0.041683 | 0.016733 | 0.040825 | 0.05164 | 0.008165 | 0.016151 | 0.01472
| Rel.Std.Dev. |  | 11.17% | 7.404% | 3.061% | 7.543% | 8.666% | 5.379% | 7.069% | 41.54% |  | 9.501% |  | 17.4% | 8.367% | 12.89% | 14.08% | 4.374% | 1.839% | 7.816%
| PDM3 |  | 195.9% | -4.759% | -19.36% | 6.951% | -18.03% | -8.662% | -8.662% | 43.64% |  | 12.42% |  | 12.22% | -6.32% | 48.33% | 71.75% | -12.57% | 311.4% | -11.78%
| Z-Score (Absolute) |  | 5.05 | 0.39 | 0.78 | 0.07 | 0.75 | 0.49 | 0.49 | 0.92 | 9.72 | 0.08 | 9.72 | 0.07 | 0.43 | 1.05 | 1.68 | 0.6 | 8.18 | 0.58
| Analytical results for Al in OREAS 290 (Certified Value 7.21 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.267 | 6.4532 | 6.79 | 7.34 | 6.82 | 6.95 | 6.86 | 5.76 | 7.41 | 7.43 | 7.24 | 7.1621 | 7.3635 | 7.21 | 6.92 | 6.95 | 7.2896 | 7.6 | 1
| 1 | 2 | 6.8019 | 6.5792 | 6.86 | 7.54 | 6.97 | 7.09 | 6.99 | 4.53 | 7.41 | 7.67 | 7.25 | 7.1885 | 7.3729 | 7.68 | 7.25 | 7.1 | 7.2737 | 7.98 | 3
| 1 | 3 | 6.9851 | 7.5206 | 7.03 | 7.63 | 6.73 | 7.43 | 7.1 | 4.79 | 7.33 | 7.48 | 7.33 | 7.4645 | 7.3065 | 7.71 | 6.97 | 6.99 | 7.2896 | 7.25 | 16
| 1 | 4 | 6.7299 | 7.0267 | 7.1 | 7.45 | 7.18 | 7.22 | 7.04 | 3.93 | 7.23 | 7.35 | 7.31 | 7.3438 | 7.3501 | 7.38 | 7.17 | 7.34 | 7.2556 | 7.19 | 7.2094
| 1 | 5 | 6.5858 | 6.2775 | 7.16 | 7.52 | 7.02 | 6.97 | 7.11 | 3.63 | 7.35 | 7.47 | 7.3 | 7.4647 | 7.4108 | 7.96 | 7.27 | 7.01 | 7.253 | 7.13 | 14
| 1 | 6 | 7.3033 | 6.7283 | 7.07 | 7.52 | 6.79 | 7.41 | 7.04 | 3.62 | 7.3 | 7.76 | 7.29 | 7.3422 | 7.4516 | 7.62 | 7.09 | 6.98 | 7.2604 | 7.3
| Mean |  | 6.9455 | 6.7643 | 7.0017 | 7.5 | 6.9183 | 7.1783 | 7.0233 | 4.3767 | 7.3383 | 7.5267 | 7.2867 | 7.3276 | 7.3759 | 7.5933 | 7.1117 | 7.0617 | 7.2703 | 7.4083
| Median |  | 6.8935 | 6.6538 | 7.05 | 7.52 | 6.895 | 7.155 | 7.04 | 4.23 | 7.34 | 7.475 | 7.295 | 7.343 | 7.3682 | 7.65 | 7.13 | 7 | 7.2671 | 7.275
| Std Dev. |  | 0.29306 | 0.44948 | 0.14497 | 0.097365 | 0.16916 | 0.21075 | 0.09136 | 0.82916 | 0.068823 | 0.15552 | 0.035024 | 0.13021 | 0.050186 | 0.2641 | 0.14483 | 0.14552 | 0.016562 | 0.32419
| Rel.Std.Dev. |  | 4.219% | 6.645% | 2.071% | 1.298% | 2.445% | 2.936% | 1.301% | 18.95% | 0.9379% | 2.066% | 0.4807% | 1.777% | 0.6804% | 3.478% | 2.037% | 2.061% | 0.2278% | 4.376%
| PDM3 |  | -3.66% | -6.174% | -2.881% | 4.031% | -4.037% | -0.4305% | -2.58% | -39.29% | 1.789% | 4.401% | 1.072% | 1.641% | 2.31% | 5.326% | -1.355% | -2.049% | 0.8454% | 2.76%
| Z-Score (Absolute) |  | 0.98 | 1.61 | 0.78 | 0.97 | 1.07 | 0.16 | 0.7 | 9.98 | 0.4 | 1.06 | 0.22 | 0.36 | 0.53 | 1.29 | 0.39 | 0.57 | 0.16 | 0.64
| Analytical results for As in OREAS 290 (Certified Value 313 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 331.38 | 268.8 | 85 | 342 | 307 | 317 | 303 | 282 | 324 | 295 | 308.3171 | 324.4 | 324 | 303 | 214 | 332.67 | 308 |  | 1
| 1 | 2 | 335.16 | 271.2 | 113.9 | 345 | 304 | 320 | 316 | 276 | 279 | 301 | 300.0561 | 322.8 | 334 | 317 | 213 | 332.925 | 317 |  | 4
| 1 | 3 | 361.62 | 285.1 | 135.8 | 344 | 290 | 328 | 317 | 300 | 322 | 300 | 293.4704 | 320.8 | 313 | 308 | 194 | 323.776 | 301 |  | 16
| 1 | 4 | 347.76 | 292.8 | 173.7 | 344 | 322 | 330 | 320 | 296 | 324 | 303 | 290.4061 | 313.6 | 304 | 312 | 184 | 332.515 | 296 |  | 313.437
| 1 | 5 | 336.42 | 271.5 | 111.4 | 322 | 311 | 300 | 314 | 280 | 311 | 303 | 299.9011 | 321.8 | 331 | 309 | 174 | 338.54 | 299 |  | 15
| 1 | 6 | 341.46 | 284.5 | 117.1 | 335 | 304 | 341 | 311 | 297 | 310 | 298 | 309.4548 | 326.9 | 328 | 302 | 193 | 330.1 | 306
| Mean |  | 342.3 | 278.9833 | 122.8167 | 338.6667 | 306.3333 | 322.6667 | 313.5 | 288.5 | 311.6667 | 300 | 300.2676 | 321.7167 | 322.3333 | 308.5 | 195.3333 | 331.7543 | 304.5
| Median |  | 338.94 | 278 | 115.5 | 343 | 305.5 | 324 | 315 | 289 | 316.5 | 300.5 | 299.9786 | 322.3 | 326 | 308.5 | 193.5 | 332.5925 | 303.5
| Std Dev. |  | 11.0325 | 9.788 | 29.7668 | 8.9368 | 10.4435 | 13.938 | 5.9582 | 10.3102 | 17.2124 | 3.0984 | 7.6528 | 4.5159 | 11.5701 | 5.6125 | 15.8198 | 4.8005 | 7.5565
| Rel.Std.Dev. |  | 3.223% | 3.508% | 24.24% | 2.639% | 3.409% | 4.32% | 1.901% | 3.574% | 5.523% | 1.033% | 2.549% | 1.404% | 3.589% | 1.819% | 8.099% | 1.447% | 2.482%
| PDM3 |  | 9.209% | -10.99% | -60.82% | 8.049% | -2.266% | 2.945% | 0.02009% | -7.956% | -0.5648% | -4.287% | -4.202% | 2.642% | 2.838% | -1.575% | -37.68% | 5.844% | -2.851%
| Z-Score (Absolute) |  | 1.65 | 1.44 | 9.05 | 1.47 | 0.11 | 0.69 | 0.24 | 0.97 | 0.15 | 0.41 | 0.4 | 0.64 | 0.67 | 0 | 5.52 | 1.13 | 0.19
| Analytical results for Ba in OREAS 290 (Certified Value 739 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 749.55 | 646 | 673 | 770 | 710 | 720 | 730 | 731 | 788 | 686 | 764 | 731.1394 | 743.5 | 783 | 714 | 760 | 738.0958 | 735 | 1
| 1 | 2 | 699.2 | 628 | 650 | 780 | 720 | 730 | 740 | 693 | 794 | 710 | 771 | 732.1881 | 716.6 | 772 | 742 | 760 | 732.9584 | 758 | 5
| 1 | 3 | 747.65 | 748 | 671 | 800 | 700 | 740 | 760 | 709 | 792 | 666 | 784 | 741.492 | 719.6 | 823 | 733 | 750 | 731.7167 | 734 | 16
| 1 | 4 | 704.9 | 619 | 675 | 780 | 740 | 730 | 750 | 657 | 783 | 708 | 787 | 732.6734 | 716.1 | 770 | 729 | 720 | 733.9598 | 754 | 738.9002
| 1 | 5 | 701.1 | 671 | 689 | 780 | 730 | 710 | 760 | 670 | 791 | 691 | 762 | 734.6008 | 727.6 | 758 | 756 | 670 | 733.6694 | 750 | 16
| 1 | 6 | 764.75 | 641 | 649 | 780 | 700 | 760 | 740 | 689 | 790 | 723 | 778 | 742.8409 | 709.3 | 769 | 741 | 680 | 730.497 | 759
| Mean |  | 727.8583 | 658.8333 | 667.8333 | 781.6667 | 716.6667 | 731.6667 | 746.6667 | 691.5 | 789.6667 | 697.3333 | 774.3333 | 735.8224 | 722.1167 | 779.1667 | 735.8333 | 723.3333 | 733.4829 | 748.3333
| Median |  | 726.275 | 643.5 | 672 | 780 | 715 | 730 | 745 | 691 | 790.5 | 699.5 | 774.5 | 733.6371 | 718.1 | 771 | 737 | 735 | 733.3139 | 752
| Std Dev. |  | 29.2835 | 47.1568 | 15.5488 | 9.8319 | 16.3299 | 17.224 | 12.1106 | 26.5612 | 3.8297 | 20.4124 | 10.3666 | 5.0587 | 12.0378 | 22.9034 | 14.1622 | 40.332 | 2.602 | 11.1833
| Rel.Std.Dev. |  | 4.023% | 7.158% | 2.328% | 1.258% | 2.279% | 2.354% | 1.622% | 3.841% | 0.485% | 2.927% | 1.339% | 0.6875% | 1.667% | 2.939% | 1.925% | 5.576% | 0.3547% | 1.494%
| PDM3 |  | -1.494% | -10.84% | -9.618% | 5.788% | -3.009% | -0.979% | 1.051% | -6.415% | 6.871% | -5.626% | 4.795% | -0.4165% | -2.271% | 5.45% | -0.4151% | -2.107% | -0.7332% | 1.277%
| Z-Score (Absolute) |  | 0.2 | 3.14 | 2.76 | 2.09 | 0.68 | 0.04 | 0.6 | 1.75 | 2.43 | 1.5 | 1.78 | 0.14 | 0.45 | 1.98 | 0.14 | 0.39 | 0.04 | 0.67
| Analytical results for Be in OREAS 290 (Certified Value 2.51 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 2.4 | 2 | 2 | 2.73 | 2.32 | 2.35 | 2.09 | 2.1 | 2.6 | 2.7 | 2.5 | 2.4134 | 2.44 | 2.4 | 2.7 | 2.65 | < 5 | 2.42 | 1
| 1 | 2 | 2.275 | 2 | 2 | 2.72 | 2.45 | 2.56 | 2.19 | 2.1 | 2.5 | 2.9 | 2.6 | 2.5227 | 2.47 | 2.4 | 2.4 | 2.78 | < 5 | 2.59 | 6
| 1 | 3 | 2.4375 | 2 | 2 | 2.86 | 2.38 | 2.56 | 2.14 | 2.4 | 2.6 | 2.8 | 2.5 | 2.5928 | 2.45 | 2.2 | 2.3 | 2.66 | < 5 | 2.57 | 16
| 1 | 4 | 2.3 | 2 | 2 | 2.77 | 2.56 | 2.52 | 2.19 | 2.1 | 2.5 | 2.8 | 2.5 | 2.3603 | 2.52 | 2.4 | 2.4 | 2.88 | < 5 | 2.62 | 2.5085
| 1 | 5 | 2.3 | 2 | 3 | 2.77 | 2.07 | 2.38 | 2.19 | 1.9 | 2.5 | 2.9 | 2.5 | 2.5673 | 2.45 | 2.5 | 2.3 | 2.48 | < 5 | 2.68 | 17
| 1 | 6 | 2.5 | 2 | 2 | 2.88 | 2.23 | 2.52 | 2.13 | 1.7 | 2.5 | 2.8 | 2.5 | 2.4684 | 2.43 | 2.5 | 2.6 | 2.81 | < 5 | 2.62
| Mean |  | 2.3687 | 2 | 2.1667 | 2.7883 | 2.335 | 2.4817 | 2.155 | 2.05 | 2.5333 | 2.8167 | 2.5167 | 2.4875 | 2.46 | 2.4 | 2.45 | 2.71 |  | 2.5833
| Median |  | 2.35 | 2 | 2 | 2.77 | 2.35 | 2.52 | 2.165 | 2.1 | 2.5 | 2.8 | 2.5 | 2.4956 | 2.45 | 2.4 | 2.4 | 2.72 |  | 2.605
| Std Dev. |  | 0.090743 | 0 | 0.40825 | 0.066758 | 0.17167 | 0.09261 | 0.041833 | 0.23452 | 0.05164 | 0.075277 | 0.040825 | 0.090266 | 0.032249 | 0.10954 | 0.16432 | 0.14339 |  | 0.088242
| Rel.Std.Dev. |  | 3.831% | 0% | 18.84% | 2.394% | 7.352% | 3.732% | 1.941% | 11.44% | 2.038% | 2.673% | 1.622% | 3.629% | 1.311% | 4.564% | 6.707% | 5.291% |  | 3.416%
| PDM3 |  | -5.571% | -20.27% | -13.63% | 11.16% | -6.916% | -1.069% | -14.09% | -18.28% | 0.9902% | 12.29% | 0.3258% | -0.8373% | -1.933% | -4.325% | -2.332% | 8.033% |  | 2.983%
| Z-Score (Absolute) |  | 0.85 | NA | NA | 2.23 | 1.1 | 0.02 | 2.42 | 3.2 | 0.36 | 2.44 | 0.24 | 0.02 | 0.18 | 0.62 | 0.25 | 1.66 | 0.11 | 0.73
Results from laboratories 2.03 and 2.04 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 290 (Certified Value 0.37 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | < 5 | 0.38 | 0.34 | 0.35 | 0.37 | 0.36 | 0.33 | 0.36 | < 5 | 0.32 | < 5 | 0.37967 | 0.39 | 0.39 | < 50 | 0.37 | < 5 | 0.39 | 1
| 1 | 2 | < 5 | 0.36 | 0.33 | 0.37 | 0.34 | 0.37 | 0.32 | 0.33 | < 5 | 0.37 | < 5 | 0.37904 | 0.39 | 0.39 | < 50 | 0.36 | < 5 | 0.39 | 7
| 1 | 3 | < 5 | 0.38 | 0.33 | 0.44 | 0.33 | 0.39 | 0.33 | 0.36 | < 5 | 0.34 | < 5 | 0.39056 | 0.43 | 0.38 | < 50 | 0.41 | < 5 | 0.34 | 16
| 1 | 4 | < 5 | 0.36 | 0.36 | 0.4 | 0.37 | 0.44 | 0.32 | 0.31 | < 5 | 0.36 | < 5 | 0.35455 | 0.38 | 0.4 | < 50 | 0.42 | < 5 | 0.4 | 0.36738
| 1 | 5 | < 5 | 0.35 | 0.34 | 0.43 | 0.36 | 0.32 | 0.32 | 0.31 | < 5 | 0.37 | < 5 | 0.38876 | 0.38 | 0.4 | < 50 | 0.37 | < 5 | 0.42 | 18
| 1 | 6 | < 5 | 0.38 | 0.33 | 0.34 | 0.37 | 0.36 | 0.31 | 0.33 | < 5 | 0.4 | < 5 | 0.34729 | 0.4 | 0.4 | < 50 | 0.37 | < 5 | 0.39
| Mean |  |  | 0.36833 | 0.33833 | 0.38833 | 0.35667 | 0.37333 | 0.32167 | 0.33333 |  | 0.36 |  | 0.37331 | 0.395 | 0.39333 |  | 0.38333 |  | 0.38833
| Median |  |  | 0.37 | 0.335 | 0.385 | 0.365 | 0.365 | 0.32 | 0.33 |  | 0.365 |  | 0.37936 | 0.39 | 0.395 |  | 0.37 |  | 0.39
| Std Dev. |  |  | 0.013292 | 0.01169 | 0.041673 | 0.017512 | 0.039833 | 0.0075277 | 0.022509 |  | 0.027568 |  | 0.018101 | 0.018708 | 0.008165 |  | 0.025033 |  | 0.026394
| Rel.Std.Dev. |  |  | 3.609% | 3.455% | 10.73% | 4.91% | 10.67% | 2.34% | 6.753% |  | 7.658% |  | 4.849% | 4.736% | 2.076% |  | 6.53% |  | 6.797%
| PDM3 |  |  | 0.2587% | -7.907% | 5.703% | -2.917% | 1.62% | -12.44% | -9.268% |  | -2.01% |  | 1.614% | 7.517% | 7.064% |  | 4.342% |  | 5.703%
| Z-Score (Absolute) |  | 61.17 | 0.43 | 1.3 | 0.14 | 0.77 | 0.29 | 1.78 | 1.44 | 61.17 | 0.67 | 61.17 | 0.29 | 0.34 | 0.29 | NA | 0 | 61.17 | 0.14
Results from laboratory 2.16 were removed due to their 10 ppm reading resolution.
| Analytical results for Ca in OREAS 290 (Certified Value 0.525 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.53742 | 0.4748 | 0.5 | 0.55 | 0.51 | 0.51 | 0.49 | 0.49 | 0.53 | 0.45 | 0.53 | 0.62356 | 0.5235 | 0.49 | 0.54 | 0.56 | 0.54191 | 0.564 | 1
| 1 | 2 | 0.50524 | 0.4717 | 0.51 | 0.56 | 0.52 | 0.52 | 0.5 | 0.45 | 0.52 | 0.46 | 0.52 | 0.64502 | 0.5239 | 0.524 | 0.57 | 0.57 | 0.548 | 0.586 | 8
| 1 | 3 | 0.53335 | 0.514 | 0.5 | 0.56 | 0.49 | 0.55 | 0.51 | 0.47 | 0.52 | 0.49 | 0.53 | 0.66303 | 0.5219 | 0.528 | 0.56 | 0.56 | 0.53478 | 0.528 | 16
| 1 | 4 | 0.51 | 0.4954 | 0.51 | 0.55 | 0.53 | 0.53 | 0.51 | 0.4 | 0.52 | 0.46 | 0.52 | 0.65407 | 0.5197 | 0.498 | 0.56 | 0.56 | 0.552 | 0.532 | 0.52455
| 1 | 5 | 0.50444 | 0.4511 | 0.52 | 0.55 | 0.52 | 0.51 | 0.51 | 0.35 | 0.52 | 0.51 | 0.52 | 0.65618 | 0.5258 | 0.538 | 0.58 | 0.56 | 0.54478 | 0.602 | 19
| 1 | 6 | 0.5473 | 0.4851 | 0.5 | 0.55 | 0.5 | 0.54 | 0.5 | 0.37 | 0.52 | 0.46 | 0.53 | 0.65883 | 0.5242 | 0.517 | 0.57 | 0.55 | 0.536 | 0.609
| Mean |  | 0.52296 | 0.48202 | 0.50667 | 0.55333 | 0.51167 | 0.52667 | 0.50333 | 0.42167 | 0.52167 | 0.47167 | 0.525 | 0.65012 | 0.52317 | 0.51583 | 0.56333 | 0.56 | 0.54291 | 0.57017
| Median |  | 0.52167 | 0.47995 | 0.505 | 0.55 | 0.515 | 0.525 | 0.505 | 0.425 | 0.52 | 0.46 | 0.525 | 0.65513 | 0.5237 | 0.5205 | 0.565 | 0.56 | 0.54334 | 0.575
| Std Dev. |  | 0.018626 | 0.021568 | 0.008165 | 0.005164 | 0.01472 | 0.01633 | 0.008165 | 0.056716 | 0.0040825 | 0.023166 | 0.0054772 | 0.014328 | 0.0021106 | 0.0184 | 0.013663 | 0.0063246 | 0.0067368 | 0.03477
| Rel.Std.Dev. |  | 3.562% | 4.475% | 1.612% | 0.9332% | 2.877% | 3.101% | 1.622% | 13.45% | 0.7826% | 4.912% | 1.043% | 2.204% | 0.4034% | 3.567% | 2.425% | 1.129% | 1.241% | 6.098%
| PDM3 |  | -0.3022% | -8.108% | -3.408% | 5.488% | -2.455% | 0.4045% | -4.044% | -19.61% | -0.5487% | -10.08% | 0.08673% | 23.94% | -0.2628% | -1.661% | 7.395% | 6.759% | 3.501% | 8.697%
| Z-Score (Absolute) |  | 0 | 1.4 | 0.56 | 1.03 | 0.39 | 0.12 | 0.67 | 3.46 | 0.05 | 1.75 | 0.07 | 4.33 | 0 | 0.25 | 1.37 | 1.26 | 0.68 | 1.61
| Analytical results for Cd in OREAS 290 (Certified Value 0.054 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | < 0.4 | 0.1 | 0.05 | 0.04 | 0.06 | 0.06 | 0.04 | 0.06 | < 1 | 0.04 | 1 | 0.078182 | 0.06 | 0.22 | < 0.2 | 0.05 | 0.595 | 0.06 | 1
| 1 | 2 | < 0.4 | 0.11 | 0.04 | 0.05 | 0.06 | 0.05 | 0.06 | 0.08 | < 1 | 0.05 | 1 | 0.065423 | 0.06 | 0.25 | < 0.2 | 0.04 | 0.60758 | 0.06 | 9
| 1 | 3 | < 0.4 | 0.11 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.09 | < 1 | 0.05 | 1 | < 0.05 | 0.05 | 0.26 | < 0.2 | < 0.02 | 0.62944 | 0.06 | 16
| 1 | 4 | < 0.4 | 0.12 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.07 | < 1 | 0.05 | < 1 | 0.057686 | 0.07 | 0.25 | < 0.2 | 0.06 | 0.61667 | 0.05 | 0.053506
| 1 | 5 | < 0.4 | 0.1 | 0.04 | 0.05 | 0.05 | 0.05 | 0.04 | 0.1 | < 1 | 0.05 | 1 | 0.050659 | 0.08 | 0.2 | < 0.2 | 0.02 | 0.61935 | 0.06 | 20
| 1 | 6 | < 0.4 | 0.11 | 0.04 | 0.05 | 0.06 | 0.05 | 0.05 | 0.08 | < 1 | 0.04 | 1 | 0.069238 | 0.05 | 0.25 | < 0.2 | 0.03 | 0.595 | 0.06
| Mean |  |  | 0.10833 | 0.043333 | 0.046667 | 0.053333 | 0.051667 | 0.046667 | 0.08 |  | 0.046667 | 1 | 0.064238 | 0.061667 | 0.23833 |  | 0.04 | 0.61051 | 0.058333
| Median |  |  | 0.11 | 0.04 | 0.05 | 0.055 | 0.05 | 0.045 | 0.08 |  | 0.05 | 1 | 0.065423 | 0.06 | 0.25 |  | 0.04 | 0.61212 | 0.06
| Std Dev. |  |  | 0.0075277 | 0.005164 | 0.005164 | 0.008165 | 0.0040825 | 0.008165 | 0.014142 |  | 0.005164 | 0 | 0.010585 | 0.01169 | 0.023166 |  | 0.015811 | 0.013888 | 0.0040825
| Rel.Std.Dev. |  |  | 6.949% | 11.92% | 11.07% | 15.31% | 7.902% | 17.5% | 17.68% |  | 11.07% | 0% | 16.48% | 18.96% | 9.72% |  | 39.53% | 2.275% | 6.999%
| PDM3 |  |  | 102.5% | -19.01% | -12.78% | -0.3235% | -3.438% | -12.78% | 49.51% |  | -12.78% | 1769% | 20.06% | 15.25% | 345.4% |  | -25.24% | 1041% | 9.021%
| Z-Score (Absolute) |  | 6.37 | 2.25 | 0.67 | 0.52 | 0.22 | 0.3 | 0.52 | 0.97 | 19.85 | 0.52 | NA | 0.03 | 0.15 | 8.09 | 1.87 | 1.05 | 24.82 | 0
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 290 (Certified Value 79 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.16 | 2.17 | 2.20 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [6 empty] | 0
1
| 1 | 1 | 63.14 | 74.27 | 87.8 | 73.9 | 81.6 | 79.8 | 69.53 | 81.6 | 81.86 | 65.5 | 77.7 | 81.7 | [6 empty] | 1
| 1 | 2 | 69.9 | 72.14 | 88.2 | 78.7 | 85.4 | 81.6 | 60.86 | 82.64 | 79.39 | 68.9 | 81.1 | 82.3 | [6 empty] | 38
| 1 | 3 | 78.62 | 74.74 | 91.3 | 72.8 | 84.5 | 81.7 | 64.91 | 83.35 | 80.68 | 66.3 | 77.6 | 79.6 | [6 empty] | 16
| 1 | 4 | 74.88 | 74.58 | 86.1 | 82.4 | 87.4 | 82 | 55.28 | 84.06 | 81.1 | 68.3 | 84.1 | 80.7 | [6 empty] | 78.7458
| 1 | 5 | 69.05 | 76.1 | 85.8 | 69.7 | 82.2 | 80.4 | 59.89 | 81.71 | 80.02 | 69.1 | 80.6 | 83.8 | [6 empty] | 21
| 1 | 6 | 68.05 | 71.29 | 86.6 | 73.4 | 86.7 | 80 | 58.84 | 85.29 | 80.16 | 69.3 | 80.8 | 81.2
| Mean |  | 70.6067 | 73.8533 | 87.6333 | 75.15 | 84.6333 | 80.9167 | 61.5517 | 83.1083 | 80.535 | 67.9 | 80.3167 | 81.55
| Median |  | 69.475 | 74.425 | 87.2 | 73.65 | 84.95 | 81 | 60.375 | 82.995 | 80.42 | 68.6 | 80.7 | 81.45
| Std Dev. |  | 5.4348 | 1.7912 | 2.0285 | 4.5829 | 2.3517 | 0.96003 | 4.9972 | 1.4264 | 0.87324 | 1.605 | 2.4277 | 1.4349
| Rel.Std.Dev. |  | 7.697% | 2.425% | 2.315% | 6.098% | 2.779% | 1.186% | 8.119% | 1.716% | 1.084% | 2.364% | 3.023% | 1.76%
| PDM3 |  | -10.34% | -6.213% | 11.29% | -4.566% | 7.477% | 2.757% | -21.83% | 5.54% | 2.272% | -13.77% | 1.995% | 3.561%
| Z-Score (Absolute) |  | 1.4 | 0.93 | 1.02 | 0.75 | 0.6 | 0.07 | 2.68 | 0.38 | 0.02 | 1.78 | 0.02 | 0.16
| Analytical results for Co in OREAS 290 (Certified Value 14.2 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 14.6416 | 14.5 | 14 | 15.1 | 13.6 | 14.1 | 13.6 | 13 | 14 | 12.3 | 12 | 14.9638 | 14.8 | 15 | 16 | 14.2 | 17.4675 | 14.9 | 1
| 1 | 2 | 13.4488 | 13.1 | 13.8 | 15.2 | 13.8 | 15 | 14.1 | 11.5 | 14 | 12.6 | 11 | 15.2691 | 14.5 | 15 | 15 | 14.7 | 17.615 | 15.1 | 10
| 1 | 3 | 14.5381 | 12.8 | 14.1 | 15.5 | 12.9 | 14.8 | 13.8 | 11.9 | 14 | 12.1 | 13 | 15.7422 | 14.3 | 15 | 14 | 13.8 | 17.76 | 13.5 | 16
| 1 | 4 | 13.6637 | 13.3 | 14.3 | 15.2 | 13.8 | 15 | 14.1 | 11 | 14 | 12.1 | 12 | 15.8435 | 14.3 | 15 | 15 | 14.4 | 17.75 | 14.8 | 14.1993
| 1 | 5 | 13.3938 | 14.4 | 14.6 | 15 | 12.3 | 13.9 | 13.8 | 11.1 | 14 | 12.7 | 12 | 15.8045 | 14.4 | 15 | 16 | 14.2 | 17.965 | 15.1 | 22
| 1 | 6 | 14.4631 | 14.2 | 14.3 | 15.6 | 12.8 | 14.5 | 13.7 | 11.6 | 14 | 12.7 | 12 | 15.5408 | 14.2 | 15 | 15 | 15.1 | 17.85 | 14.3
| Mean |  | 14.0249 | 13.7167 | 14.1833 | 15.2667 | 13.2 | 14.55 | 13.85 | 11.6833 | 14 | 12.4167 | 12 | 15.5273 | 14.4167 | 15 | 15.1667 | 14.4 | 17.7346 | 14.6167
| Median |  | 14.0634 | 13.75 | 14.2 | 15.2 | 13.25 | 14.65 | 13.8 | 11.55 | 14 | 12.45 | 12 | 15.6415 | 14.35 | 15 | 15 | 14.3 | 17.755 | 14.85
| Std Dev. |  | 0.58247 | 0.73598 | 0.27869 | 0.23381 | 0.6229 | 0.46797 | 0.20736 | 0.72503 | 0 | 0.28577 | 0.63246 | 0.34854 | 0.2137 | 0 | 0.75277 | 0.45166 | 0.17489 | 0.62102
| Rel.Std.Dev. |  | 4.153% | 5.366% | 1.965% | 1.532% | 4.719% | 3.216% | 1.497% | 6.206% | 0% | 2.302% | 5.27% | 2.245% | 1.482% | 0% | 4.963% | 3.137% | 0.9861% | 4.249%
| PDM3 |  | -1.229% | -3.399% | -0.1128% | 7.517% | -7.038% | 2.47% | -2.46% | -17.72% | -1.404% | -12.55% | -15.49% | 9.352% | 1.53% | 5.639% | 6.812% | 1.413% | 24.9% | 2.939%
| Z-Score (Absolute) |  | 0.35 | 0.76 | 0.14 | 1.29 | 1.45 | 0.34 | 0.59 | 3.46 | NA | 2.49 | NA | 1.64 | 0.17 | NA | NA | 0.14 | 4.57 | 0.43
Results from laboratories 2.10, 2.12, 2.15 and 2.16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 290 (Certified Value 123 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 100.0211 | 121 | 116 | 130 | 116 | 122 | 124 | 51 | 124 | 122 | 133 | 123.6643 | 124 | 99 | 134 | 103 | 118.12 | 120 | 1
| 1 | 2 | 90.7555 | 123 | 114 | 131 | 116 | 123 | 125 | 41 | 118 | 120 | 130 | 127.5592 | 124 | 101 | 136 | 99 | 120.325 | 120 | 11
| 1 | 3 | 99.4454 | 122 | 113 | 136 | 112 | 125 | 126 | 42 | 125 | 119 | 123 | 115.3711 | 125 | 104 | 134 | 93 | 119.88 | 133 | 16
| 1 | 4 | 91.541 | 127 | 119 | 132 | 125 | 122 | 127 | 47 | 127 | 121 | 128 | 115.3585 | 122 | 110 | 130 | 87 | 118.5842 | 127 | 122.7493
| 1 | 5 | 89.5206 | 121 | 122 | 132 | 117 | 118 | 126 | 46 | 129 | 124 | 131 | 122.4501 | 126 | 116 | 138 | 111 | 118.27 | 117 | 23
| 1 | 6 | 98.4024 | 117 | 118 | 134 | 114 | 125 | 127 | 46 | 128 | 125 | 117 | 126.4731 | 122 | 121 | 133 | 84 | 117.985 | 124
| Mean |  | 94.9477 | 121.8333 | 117 | 132.5 | 116.6667 | 122.5 | 125.8333 | 45.5 | 125.1667 | 121.8333 | 127 | 121.8127 | 123.8333 | 108.5 | 134.1667 | 96.1667 | 118.8607 | 123.5
| Median |  | 94.9717 | 121.5 | 117 | 132 | 116 | 122.5 | 126 | 46 | 126 | 121.5 | 129 | 123.0572 | 124 | 107 | 134 | 96 | 118.4271 | 122
| Std Dev. |  | 4.8278 | 3.2506 | 3.3466 | 2.1679 | 4.4572 | 2.5884 | 1.169 | 3.6194 | 3.9707 | 2.3166 | 5.9666 | 5.324 | 1.6021 | 8.735 | 2.7142 | 10.1669 | 0.99237 | 5.8224
| Rel.Std.Dev. |  | 5.085% | 2.668% | 2.86% | 1.636% | 3.82% | 2.113% | 0.929% | 7.955% | 3.172% | 1.901% | 4.698% | 4.371% | 1.294% | 8.051% | 2.023% | 10.57% | 0.8349% | 4.714%
| PDM3 |  | -22.65% | -0.7462% | -4.684% | 7.944% | -4.955% | -0.2031% | 2.512% | -62.93% | 1.969% | -0.7462% | 3.463% | -0.763% | 0.8831% | -11.61% | 9.301% | -21.66% | -3.168% | 0.6115%
| Z-Score (Absolute) |  | 4.1 | 0 | 0.74 | 1.63 | 0.79 | 0.1 | 0.61 | 11.65 | 0.51 | 0 | 0.79 | 0 | 0.31 | 2.04 | 1.88 | 3.92 | 0.45 | 0.25
| Analytical results for Cs in OREAS 290 (Certified Value 10.3 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 8.7 | 9.6 | 10.95 | 9.28 | 10.55 | 10.5 | 9 | 10.7 | 10.9068 | 10.53 | 10.4 | 10.8 | 10.3 | 9.8 |  |  |  |  | 1
| 1 | 2 | 9.5 | 9.3 | 10.85 | 9.77 | 10.95 | 10.8 | 9 | 10.8 | 10.9804 | 10.22 | 11 | 10.4 | 10.2 | 10.1 |  |  |  |  | 41
| 1 | 3 | 9.9 | 9.8 | 10.95 | 9.15 | 11.1 | 10.55 | 10 | 10.8 | 11.0048 | 10.09 | 10.9 | 10.1 | 10.1 | 9.1 |  |  |  |  | 16
| 1 | 4 | 10.1 | 9.7 | 10.85 | 9.92 | 11.2 | 10.65 | 9 | 10.9 | 11.0113 | 10.15 | 10.4 | 9.9 | 10.5 | 10.1 |  |  |  |  | 10.33
| 1 | 5 | 8.6 | 10 | 10.65 | 8.74 | 10.4 | 10.2 | 9 | 11.4 | 10.8825 | 10.32 | 10.7 | 10.2 | 9.96 | 10.2 |  |  |  |  | 24
| 1 | 6 | 10.6 | 9.6 | 10.7 | 9.17 | 10.85 | 10.3 | 9 | 11.2 | 10.9616 | 10.23 | 11 | 11.2 | 9.92 | 10
| Mean |  | 9.5667 | 9.6667 | 10.825 | 9.3383 | 10.8417 | 10.5 | 9.1667 | 10.9667 | 10.9579 | 10.2567 | 10.7333 | 10.4333 | 10.1633 | 9.8833
| Median |  | 9.7 | 9.65 | 10.85 | 9.225 | 10.9 | 10.525 | 9 | 10.85 | 10.971 | 10.225 | 10.8 | 10.3 | 10.15 | 10.05
| Std Dev. |  | 0.79415 | 0.23381 | 0.1255 | 0.43586 | 0.31212 | 0.22136 | 0.40825 | 0.27325 | 0.052681 | 0.15488 | 0.28048 | 0.48442 | 0.21814 | 0.40702
| Rel.Std.Dev. |  | 8.301% | 2.419% | 1.159% | 4.667% | 2.879% | 2.108% | 4.454% | 2.492% | 0.4808% | 1.51% | 2.613% | 4.643% | 2.146% | 4.118%
| PDM3 |  | -7.389% | -6.421% | 4.792% | -9.6% | 4.954% | 1.646% | -11.26% | 6.164% | 6.079% | -0.7096% | 3.905% | 1.001% | -1.613% | -4.324%
| Z-Score (Absolute) |  | 1.43 | 1.27 | 0.65 | 1.81 | 0.67 | 0.11 | NA | 0.88 | 0.87 | 0.29 | 0.5 | 0 | 0.45 | 0.91
| Analytical results for Cu in OREAS 290 (Certified Value 32.6 ppm) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 34.2789 | 31.6 | 31.4 | 31.2 | 30.6 | 32.4 | 31.2 | 37 | 26.1 | 32.1 | 30.8 | 33.9839 | 34.8 | 32 | 32 | 32.8 | 35.0591 | 36 | 1
| 1 | 2 | 33.5854 | 31.1 | 30.2 | 31.3 | 32.4 | 33.6 | 31.8 | 32.2 | 28.9 | 32.3 | 30.6 | 34.9355 | 33.8 | 32 | 37 | 32.2 | 35.2298 | 36.5 | 12
| 1 | 3 | 34.463 | 30.5 | 31.3 | 32.3 | 30.8 | 34.4 | 31.9 | 31.7 | 28.8 | 32 | 30.5 | 34.6696 | 33.2 | 33 | 35 | 36.2 | 34.6583 | 35.2 | 16
| 1 | 4 | 33.9688 | 30.9 | 31.5 | 31.1 | 33.9 | 35.6 | 32.6 | 30.8 | 28.4 | 32.9 | 30.8 | 33.2502 | 34.4 | 27 | 33 | 45.9 | 35.43 | 33.6 | 32.5699
| 1 | 5 | 32.5933 | 30.5 | 33.1 | 31.1 | 28.3 | 33.4 | 32.2 | 29.9 | 29 | 33.1 | 30.5 | 33.5179 | 33.9 | 32 | 35 | 34.9 | 34.8864 | 33.6 | 25
| 1 | 6 | 33.8325 | 31 | 31.9 | 31.9 | 30.2 | 36 | 31.3 | 31.7 | 28.2 | 33.9 | 30.8 | 33.3638 | 34.9 | 31 | 33 | 33.8 | 35.54 | 33.2
| Mean |  | 33.787 | 30.9333 | 31.5667 | 31.4833 | 31.0333 | 34.2333 | 31.8333 | 32.2167 | 28.2333 | 32.7167 | 30.6667 | 33.9535 | 34.1667 | 31.1667 | 34.1667 | 35.9667 | 35.1339 | 34.6833
| Median |  | 33.9006 | 30.95 | 31.45 | 31.25 | 30.7 | 34 | 31.85 | 31.7 | 28.6 | 32.6 | 30.7 | 33.7509 | 34.15 | 32 | 34 | 34.35 | 35.1445 | 34.4
| Std Dev. |  | 0.66327 | 0.41312 | 0.94163 | 0.49967 | 1.9232 | 1.3765 | 0.53166 | 2.4815 | 1.0893 | 0.72778 | 0.15055 | 0.70855 | 0.6532 | 2.137 | 1.8348 | 5.0749 | 0.33311 | 1.4034
| Rel.Std.Dev. |  | 1.963% | 1.336% | 2.983% | 1.587% | 6.197% | 4.021% | 1.67% | 7.702% | 3.858% | 2.224% | 0.4909% | 2.087% | 1.912% | 6.857% | 5.37% | 14.11% | 0.9481% | 4.046%
| PDM3 |  | 3.737% | -5.025% | -3.08% | -3.336% | -4.718% | 5.107% | -2.261% | -1.084% | -13.31% | 0.4507% | -5.843% | 4.248% | 4.903% | -4.308% | 4.903% | 10.43% | 7.873% | 6.489%
| Z-Score (Absolute) |  | 0.59 | 0.68 | 0.4 | 0.44 | 0.64 | 0.79 | 0.28 | 0.11 | 1.89 | 0.11 | 0.8 | 0.66 | 0.76 | 0.58 | 0.76 | 1.56 | 1.19 | 0.99
| Analytical results for Dy in OREAS 290 (Certified Value 3.01 ppm) | [19 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
2
| 1 | 1 | 2.7 | 2.7 | 3.51 | 2.5 | 3.31 | [13 empty] | 1
| 1 | 2 | 3.1 | 2.7 | 3.4 | 2.7 | 3.27 | [13 empty] | 18
| 1 | 3 | 3.3 | 2.7 | 3.53 | 2.5 | 3 | [13 empty] | 16
| 1 | 4 | 3.3 | 2.7 | 3.52 | 2.6 | 3.1 | [13 empty] | 3.0127
| 1 | 5 | 2.8 | 2.8 | 3.46 | 2.7 | 3.33 | [13 empty] | 26
| 1 | 6 | 3.2 | 2.6 | 3.44 | 2.7 | 3.21
| Mean |  | 3.0667 | 2.7 | 3.4767 | 2.6167 | 3.2033
| Median |  | 3.15 | 2.7 | 3.485 | 2.65 | 3.24
| Std Dev. |  | 0.2582 | 0.063246 | 0.05164 | 0.098319 | 0.12956
| Rel.Std.Dev. |  | 8.42% | 2.342% | 1.485% | 3.757% | 4.045%
| PDM3 |  | 1.792% | -10.38% | 15.4% | -13.14% | 6.329%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.75 | 0.83 | 0.25
| Analytical results for Er in OREAS 290 (Certified Value 1.6 ppm) | [19 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
2
| 1 | 1 | 1.4 | 1.4 | 1.8 | 1.5 | 1.7 | [13 empty] | 1
| 1 | 2 | 1.4 | 1.4 | 1.93 | 1.6 | 1.69 | [13 empty] | 19
| 1 | 3 | 1.7 | 1.4 | 2.37 | 1.6 | 1.57 | [13 empty] | 16
| 1 | 4 | 1.6 | 1.3 | 1.94 | 1.5 | 1.62 | [13 empty] | 1.6002
| 1 | 5 | 1.5 | 1.4 | 1.87 | 1.6 | 1.71 | [13 empty] | 27
| 1 | 6 | 1.6 | 1.3 | 1.84 | 1.6 | 1.66
| Mean |  | 1.5333 | 1.3667 | 1.9583 | 1.5667 | 1.6583
| Median |  | 1.55 | 1.4 | 1.9 | 1.6 | 1.675
| Std Dev. |  | 0.12111 | 0.05164 | 0.20856 | 0.05164 | 0.054191
| Rel.Std.Dev. |  | 7.898% | 3.779% | 10.65% | 3.296% | 3.268%
| PDM3 |  | -4.179% | -14.59% | 22.38% | -2.096% | 3.633%
| Z-Score (Absolute) |  | 0.25 | 1.47 | 2.88 | 0 | 0.67
| Analytical results for Eu in OREAS 290 (Certified Value 1.14 ppm) | [19 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 1 | 1 | 1.23 | 1.2 | 1.18 | [13 empty] | 1
| 1 | 2 | 1.1 | 1 | 1.21 | 1.2 | 1.23 | [13 empty] | 20
| 1 | 3 | 1.3 | 1 | 1.2 | 1.1 | 1.12 | [13 empty] | 16
| 1 | 4 | 1.2 | 1 | 1.22 | 1.2 | 1.21 | [13 empty] | 1.141
| 1 | 5 | 1 | 1 | 1.2 | 1.1 | 1.21 | [13 empty] | 28
| 1 | 6 | 1.2 | 1 | 1.27 | 1.2 | 1.15
| Mean |  | 1.1333 | 1 | 1.2217 | 1.1667 | 1.1833
| Median |  | 1.15 | 1 | 1.215 | 1.2 | 1.195
| Std Dev. |  | 0.12111 | 0 | 0.026394 | 0.05164 | 0.041793
| Rel.Std.Dev. |  | 10.69% | 0% | 2.161% | 4.426% | 3.532%
| PDM3 |  | -0.6719% | -12.36% | 7.07% | 2.249% | 3.71%
| Z-Score (Absolute) |  | 0.67 | 3.37 | 1.11 | 0 | 0.34
| Analytical results for Fe in OREAS 290 (Certified Value 3.79 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.9871 | 3.7471 | 3.66 | 3.97 | 3.64 | 3.61 | 3.53 | 3.77 | 3.9 | 3.5 | 3.84 | 3.8829 | 3.95 | 3.88 | 3.75 | 3.79 | 3.63 | 3.83 | 1
| 1 | 2 | 3.7374 | 3.7825 | 3.67 | 4 | 3.66 | 3.67 | 3.59 | 3.62 | 3.86 | 3.48 | 3.81 | 3.9119 | 3.94 | 4.36 | 3.91 | 3.78 | 3.65 | 3.96 | 13
| 1 | 3 | 4.0037 | 3.8199 | 3.67 | 4.09 | 3.55 | 3.79 | 3.65 | 3.66 | 3.87 | 3.43 | 3.84 | 4.0257 | 3.94 | 4.32 | 3.84 | 3.77 | 3.6417 | 3.57 | 16
| 1 | 4 | 3.7574 | 3.8532 | 3.73 | 3.99 | 3.8 | 3.7 | 3.62 | 3.54 | 3.8 | 3.47 | 3.82 | 3.9661 | 3.91 | 3.91 | 3.83 | 3.74 | 3.59 | 3.44 | 3.7892
| 1 | 5 | 3.793 | 3.6577 | 3.83 | 3.99 | 3.7 | 3.6 | 3.64 | 3.46 | 3.86 | 3.44 | 3.85 | 4.0313 | 3.95 | 4.49 | 3.99 | 3.57 | 3.655 | 4.02 | 29
| 1 | 6 | 4.0939 | 3.7929 | 3.75 | 4.03 | 3.61 | 3.82 | 3.6 | 3.54 | 3.83 | 3.5 | 3.85 | 3.9715 | 3.97 | 4.29 | 3.9 | 3.58 | 3.6219 | 4.09
| Mean |  | 3.8954 | 3.7755 | 3.7183 | 4.0117 | 3.66 | 3.6983 | 3.605 | 3.5983 | 3.8533 | 3.47 | 3.835 | 3.9649 | 3.9433 | 4.2083 | 3.87 | 3.705 | 3.6314 | 3.8183
| Median |  | 3.8901 | 3.7877 | 3.7 | 3.995 | 3.65 | 3.685 | 3.61 | 3.58 | 3.86 | 3.475 | 3.84 | 3.9688 | 3.945 | 4.305 | 3.87 | 3.755 | 3.6358 | 3.895
| Std Dev. |  | 0.15102 | 0.067888 | 0.065853 | 0.043089 | 0.085088 | 0.091086 | 0.043243 | 0.10926 | 0.034448 | 0.029665 | 0.016432 | 0.059472 | 0.019664 | 0.2523 | 0.082219 | 0.10213 | 0.023722 | 0.26057
| Rel.Std.Dev. |  | 3.877% | 1.798% | 1.771% | 1.074% | 2.325% | 2.463% | 1.2% | 3.036% | 0.894% | 0.8549% | 0.4285% | 1.5% | 0.4987% | 5.995% | 2.125% | 2.756% | 0.6532% | 6.824%
| PDM3 |  | 2.803% | -0.3608% | -1.871% | 5.871% | -3.41% | -2.399% | -4.862% | -5.038% | 1.692% | -8.424% | 1.208% | 4.636% | 4.067% | 11.06% | 2.132% | -2.223% | -4.164% | 0.7683%
| Z-Score (Absolute) |  | 0.56 | 0.12 | 0.45 | 1.23 | 0.78 | 0.56 | 1.1 | 1.14 | 0.32 | 1.87 | 0.22 | 0.96 | 0.84 | 2.36 | 0.42 | 0.53 | 0.95 | 0.12
| Analytical results for Ga in OREAS 290 (Certified Value 18.6 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 19.01 | 17.07 | 19.95 | 17.9 | 18.6 | 17.7 | 15.8 | 17.5 | 18.6548 | 18.59 | 19.8 | 18.8 | 20.1 | 14.6506 | 19.5 |  |  |  | 1
| 1 | 2 | 17.48 | 16.96 | 20 | 19.45 | 19.45 | 18.15 | 16.2 | 18.6 | 18.5593 | 18.59 | 20.5 | 19.6 | 20.6 | 14.6249 | 19.4 |  |  |  | 46
| 1 | 3 | 18.27 | 17.14 | 20.5 | 18.3 | 19.45 | 18.1 | 16.2 | 18.6 | 18.2394 | 18.2 | 20.5 | 18.6 | 19.1 | 14.73 | 19.2 |  |  |  | 16
| 1 | 4 | 18.18 | 17.55 | 19.9 | 18 | 19.95 | 18.45 | 16.2 | 17.7 | 18.2957 | 18.63 | 18.8 | 18.8 | 20.1 | 14.5238 | 19.3 |  |  |  | 18.6282
| 1 | 5 | 19.66 | 17.89 | 19.5 | 16.6 | 18.45 | 18.1 | 15.7 | 18.4 | 18.7496 | 18.18 | 19.2 | 19.5 | 19.8 | 14.6325 | 19.7 |  |  |  | 30
| 1 | 6 | 18.23 | 17.34 | 19.85 | 17.85 | 19.25 | 17.95 | 16.4 | 19.7 | 18.1277 | 18.17 | 20.3 | 19.3 | 21.2 | 14.43 | 18.9
| Mean |  | 18.4717 | 17.325 | 19.95 | 18.0167 | 19.1917 | 18.075 | 16.0833 | 18.4167 | 18.4377 | 18.3933 | 19.85 | 19.1 | 20.15 | 14.5986 | 19.3333
| Median |  | 18.25 | 17.24 | 19.925 | 17.95 | 19.35 | 18.1 | 16.2 | 18.5 | 18.4275 | 18.395 | 20.05 | 19.05 | 20.1 | 14.6287 | 19.35
| Std Dev. |  | 0.75756 | 0.34715 | 0.32249 | 0.91469 | 0.56782 | 0.24648 | 0.27142 | 0.78337 | 0.2509 | 0.23071 | 0.71764 | 0.41952 | 0.71204 | 0.10568 | 0.27325
| Rel.Std.Dev. |  | 4.101% | 2.004% | 1.616% | 5.077% | 2.959% | 1.364% | 1.688% | 4.254% | 1.361% | 1.254% | 3.615% | 2.196% | 3.534% | 0.7239% | 1.413%
| PDM3 |  | -0.8402% | -6.996% | 7.096% | -3.283% | 3.025% | -2.97% | -13.66% | -1.135% | -1.022% | -1.261% | 6.559% | 2.533% | 8.17% | -21.63% | 3.785%
| Z-Score (Absolute) |  | 0.03 | 1 | 1.35 | 0.38 | 0.67 | 0.32 | 2.11 | 0.02 | 0 | 0.04 | 1.26 | 0.59 | 1.53 | 3.43 | 0.8
| Analytical results for Gd in OREAS 290 (Certified Value 4.47 ppm) | [19 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 3.7 | 4.2 | 5 | 4 | 4.78 | [13 empty] | 1
| 1 | 2 | 4.4 | 4 | 4.8 | 3.7 | 4.86 | [13 empty] | 21
| 1 | 3 | 5.3 | 4.1 | 4.98 | 4.2 | 4.45 | [13 empty] | 16
| 1 | 4 | 4.8 | 4.1 | 4.9 | 4 | 4.81 | [13 empty] | 4.4747
| 1 | 5 | 4.2 | 4.2 | 4.8 | 4.3 | 5 | [13 empty] | 31
| 1 | 6 | 4.9 | 4 | 4.81 | 4.2 | 4.75
| Mean |  | 4.55 | 4.1 | 4.8817 | 4.0667 | 4.775
| Median |  | 4.6 | 4.1 | 4.855 | 4.1 | 4.795
| Std Dev. |  | 0.56833 | 0.089443 | 0.092177 | 0.21602 | 0.18185
| Rel.Std.Dev. |  | 12.49% | 2.182% | 1.888% | 5.312% | 3.808%
| PDM3 |  | 1.684% | -8.373% | 9.096% | -9.118% | 6.712%
| Z-Score (Absolute) |  | 0 | 0.91 | 0.67 | 0.98 | 0.46
| Analytical results for Ge in OREAS 290 (Certified Value 0.22 ppm) | [19 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 0.63 | 0.24 | 0.11 | 0.29 | 0.25 | < 0.1 | 0.5 | 1.1688 | 1.6 | 1.8 | < 0.5 | 0.22 | 2.26 | [5 empty] | 1
| 1 | 2 | 0.85 | 0.28 | 0.12 | 0.25 | 0.26 | < 0.1 | 0.6 | 1.2101 | 1.4 | 2 | < 0.5 | 0.22 | 2.23 | [5 empty] | 48
| 1 | 3 | 0.5 | 0.24 | 0.12 | 0.22 | 0.27 | < 0.1 | 0.6 | 1.102 | 1.6 | 1.8 | < 0.5 | 0.2 | 2.15 | [5 empty] | 16
| 1 | 4 | 0.88 | 0.25 | 0.12 | 0.24 | 0.28 | < 0.1 | 0.4 | 1.1665 | 1.5 | 1.8 | < 0.5 | 0.19 | 2.11 | [5 empty] | 0.2172
| 1 | 5 | 0.83 | 0.3 | 0.1 | 0.24 | 0.27 | < 0.1 | 0.4 | 1.1545 | 1.4 | 1.9 | < 0.5 | 0.21 | 2.28 | [5 empty] | 32
| 1 | 6 | 0.7 | 0.26 | 0.11 | 0.23 | 0.28 | < 0.1 | 0.6 | 1.1347 | 1.4 | 1.9 | < 0.5 | 0.2 | 2.35
| Mean |  | 0.73167 | 0.26167 | 0.11333 | 0.245 | 0.26833 |  | 0.51667 | 1.1561 | 1.4833 | 1.8667 |  | 0.20667 | 2.23
| Median |  | 0.765 | 0.255 | 0.115 | 0.24 | 0.27 |  | 0.55 | 1.1605 | 1.45 | 1.85 |  | 0.205 | 2.245
| Std Dev. |  | 0.14878 | 0.024014 | 0.008165 | 0.02429 | 0.01169 |  | 0.098319 | 0.03625 | 0.098319 | 0.08165 |  | 0.012111 | 0.087864
| Rel.Std.Dev. |  | 20.33% | 9.177% | 7.204% | 9.914% | 4.357% |  | 19.03% | 3.136% | 6.628% | 4.374% |  | 5.86% | 3.94%
| PDM3 |  | 236.9% | 20.47% | -47.82% | 12.8% | 23.54% |  | 137.9% | 432.3% | 582.9% | 759.4% |  | -4.85% | 926.7%
| Z-Score (Absolute) |  | 3.35 | 0.04 | 1 | 0.08 | 0.09 | 1.45 | NA | 6.33 | NA | NA | 0.04 | 0.35 | 13.89
Results from laboratories 2.11, 2.14 and 2.15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 290 (Certified Value 3.98 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 3.97 | 3.32 | 4.1 | 4.2 | 4 | 3.6 | 3.67 | 4.14 | 4.0047 | 3.93 | 4.3 | 3.6 | 4.1 | 4.28 |  |  |  |  | 1
| 1 | 2 | 4.38 | 3.23 | 4 | 4.1 | 4 | 3.8 | 3.74 | 4.35 | 3.878 | 3.99 | 4.3 | 3.3 | 4.1 | 4.23 |  |  |  |  | 32
| 1 | 3 | 4.43 | 3.36 | 4.1 | 3.8 | 4 | 3.8 | 3.73 | 4.26 | 3.8019 | 4.15 | 4 | 3.5 | 4.1 | 4.27 |  |  |  |  | 16
| 1 | 4 | 4.41 | 3.35 | 4 | 3.9 | 4 | 3.8 | 3.64 | 4.37 | 3.8761 | 4.09 | 3.9 | 3.5 | 3.7 | 4.18 |  |  |  |  | 3.9792
| 1 | 5 | 4.3 | 3.39 | 4.1 | 3.7 | 3.7 | 3.7 | 3.6 | 4.33 | 3.999 | 4.04 | 4 | 3.5 | 3.8 | 4.06 |  |  |  |  | 33
| 1 | 6 | 4.44 | 3.24 | 4 | 3.9 | 4.1 | 3.7 | 3.63 | 4.32 | 3.8437 | 3.95 | 4 | 3.5 | 3.8 | 4.07
| Mean |  | 4.3217 | 3.315 | 4.05 | 3.9333 | 3.9667 | 3.7333 | 3.6683 | 4.295 | 3.9006 | 4.025 | 4.0833 | 3.4833 | 3.9333 | 4.1817
| Median |  | 4.395 | 3.335 | 4.05 | 3.9 | 4 | 3.75 | 3.655 | 4.325 | 3.8771 | 4.015 | 4 | 3.5 | 3.95 | 4.205
| Std Dev. |  | 0.17949 | 0.065955 | 0.054772 | 0.18619 | 0.13663 | 0.08165 | 0.056362 | 0.084558 | 0.083183 | 0.084794 | 0.17224 | 0.098319 | 0.18619 | 0.09704
| Rel.Std.Dev. |  | 4.153% | 1.99% | 1.352% | 4.734% | 3.444% | 2.187% | 1.536% | 1.969% | 2.133% | 2.107% | 4.218% | 2.823% | 4.734% | 2.321%
| PDM3 |  | 8.605% | -16.69% | 1.778% | -1.154% | -0.3162% | -6.18% | -7.813% | 7.935% | -1.978% | 1.15% | 2.616% | -12.46% | -1.154% | 5.087%
| Z-Score (Absolute) |  | 1.43 | 2.45 | 0.39 | 0.06 | 0.06 | 0.83 | 1.09 | 1.33 | 0.19 | 0.29 | 0.51 | 1.8 | 0.06 | 0.89
| Analytical results for Hg in OREAS 290 (Indicative Value 0.11 ppm) | [19 empty] | Indicative
Hg | Round | Replicate | Lab | [17 empty] | 1
| No | No | 2.20 | [17 empty] | ppm
|  |  | 4A*MS | [17 empty] | 2
2
| 1 | 1 | 0.11 | [17 empty] | 1
| 1 | 2 | 0.1 | [17 empty] | 15
| 1 | 3 | 0.1 | [17 empty] | 16
| 1 | 4 | 0.09 | [17 empty] | 0.10667
| 1 | 5 | 0.12 | [17 empty] | 21
| 1 | 6 | 0.12
| Mean |  | 0.10667
| Median |  | 0.105
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 11.35%
| PDM3 |  | -3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 290 (Certified Value 0.56 ppm) | [19 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
2
| 1 | 1 | 0.5 | 0.5 | 0.63 | 0.6 | 0.6 | [13 empty] | 1
| 1 | 2 | 0.5 | 0.5 | 0.66 | 0.4 | 0.6 | [13 empty] | 17
| 1 | 3 | 0.6 | 0.5 | 0.67 | 0.5 | 0.57 | [13 empty] | 16
| 1 | 4 | 0.6 | 0.5 | 0.66 | 0.5 | 0.57 | [13 empty] | 0.56167
| 1 | 5 | 0.6 | 0.5 | 0.76 | 0.5 | 0.61 | [13 empty] | 34
| 1 | 6 | 0.6 | 0.5 | 0.63 | 0.5 | 0.6
| Mean |  | 0.56667 | 0.5 | 0.66833 | 0.5 | 0.59167
| Median |  | 0.6 | 0.5 | 0.66 | 0.5 | 0.6
| Std Dev. |  | 0.05164 | 0 | 0.047924 | 0.063246 | 0.017224
| Rel.Std.Dev. |  | 9.113% | 0% | 7.171% | 12.65% | 2.911%
| PDM3 |  | 0.8902% | -10.98% | 18.99% | -10.98% | 5.341%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.03 | 0.67 | 0.25
| Analytical results for In in OREAS 290 (Certified Value 0.064 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.07 | 0.06 | 0.072 | 0.061 | 0.061 | 0.064 | 0.07 | 0.06 | 0.063966 | 0.07 | 0.05 | 0.06 | 0.057 | < 2 | 0.066 |  |  |  | 1
| 1 | 2 | 0.07 | 0.06 | 0.074 | 0.059 | 0.061 | 0.072 | 0.06 | 0.07 | 0.062995 | 0.07 | < 0.05 | 0.09 | 0.06 | < 2 | 0.07 |  |  |  | 34
| 1 | 3 | 0.08 | 0.05 | 0.069 | 0.06 | 0.062 | 0.066 | 0.07 | 0.06 | 0.063066 | 0.07 | 0.05 | 0.09 | 0.056 | < 2 | 0.059 |  |  |  | 16
| 1 | 4 | 0.07 | 0.06 | 0.074 | 0.067 | 0.066 | 0.067 | 0.06 | 0.06 | 0.065122 | 0.08 | 0.08 | 0.05 | 0.066 | < 2 | 0.067 |  |  |  | 0.064435
| 1 | 5 | 0.07 | 0.06 | 0.075 | 0.057 | 0.066 | 0.069 | 0.06 | 0.07 | 0.064363 | 0.07 | < 0.05 | 0.07 | 0.059 | < 2 | 0.068 |  |  |  | 35
| 1 | 6 | 0.08 | 0.06 | 0.076 | 0.057 | 0.068 | 0.063 | 0.06 | 0.07 | 0.065015 | 0.06 | < 0.05 | 0.06 | 0.058 | < 2 | 0.064
| Mean |  | 0.073333 | 0.058333 | 0.073333 | 0.060167 | 0.064 | 0.066833 | 0.063333 | 0.065 | 0.064088 | 0.07 | 0.06 | 0.07 | 0.059333 |  | 0.065667
| Median |  | 0.07 | 0.06 | 0.074 | 0.0595 | 0.064 | 0.0665 | 0.06 | 0.065 | 0.064164 | 0.07 | 0.05 | 0.065 | 0.0585 |  | 0.0665
| Std Dev. |  | 0.005164 | 0.0040825 | 0.0025033 | 0.0037103 | 0.0030332 | 0.0033116 | 0.005164 | 0.0054772 | 0.00092274 | 0.0063246 | 0.017321 | 0.016733 | 0.003559 |  | 0.0038297
| Rel.Std.Dev. |  | 7.042% | 6.999% | 3.414% | 6.167% | 4.739% | 4.955% | 8.154% | 8.427% | 1.44% | 9.035% | 28.87% | 23.9% | 5.998% |  | 5.832%
| PDM3 |  | 13.81% | -9.469% | 13.81% | -6.624% | -0.6748% | 3.722% | -1.709% | 0.8771% | -0.5388% | 8.637% | -6.883% | 8.637% | -7.917% |  | 1.912%
| Z-Score (Absolute) |  | 1.12 | 0.9 | 1.12 | 0.65 | 0.13 | 0.25 | 0.22 | 0 | 0.12 | 0.67 | 3.03 | 0.67 | 0.76 | 126.1 | 0.09
| Analytical results for K in OREAS 290 (Certified Value 2.8 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.332 | 2.733 | 2.72 | 2.85 | 2.7 | 2.71 | 2.68 | 2.81 | 2.97 | 2.34 | 2.73 | 3.2734 | 2.7925 | 2.68 | 2.83 | 2.78 | 2.649 | 2.98 | 1
| 1 | 2 | 2.179 | 2.779 | 2.72 | 2.97 | 2.74 | 2.75 | 2.77 | 2.83 | 2.87 | 2.29 | 2.73 | 3.3029 | 2.7875 | 2.85 | 2.92 | 2.81 | 2.6584 | 3.03 | 4
| 1 | 3 | 2.3574 | 2.574 | 2.77 | 2.98 | 2.66 | 2.82 | 2.76 | 2.85 | 2.91 | 2.28 | 2.77 | 3.4354 | 2.7522 | 2.86 | 2.9 | 2.8 | 2.6539 | 2.84 | 16
| 1 | 4 | 2.6479 | 2.786 | 2.79 | 2.93 | 2.8 | 2.78 | 2.72 | 2.79 | 2.9 | 2.34 | 2.75 | 3.4699 | 2.7714 | 2.73 | 2.88 | 2.09 | 2.649 | 2.72 | 2.8003
| 1 | 5 | 2.166 | 2.688 | 2.84 | 2.94 | 2.76 | 2.69 | 2.66 | 2.76 | 2.94 | 2.28 | 2.77 | 3.4209 | 2.8014 | 2.95 | 2.97 | 1.73 | 2.71 | 3.16 | 36
| 1 | 6 | 2.6019 | 2.72 | 2.82 | 2.95 | 2.69 | 2.86 | 2.74 | 2.84 | 2.95 | 2.31 | 2.76 | 3.4366 | 2.7989 | 2.83 | 2.92 | 2.01 | 2.645 | 3.2
| Mean |  | 2.3807 | 2.7133 | 2.7767 | 2.9367 | 2.725 | 2.7683 | 2.7217 | 2.8133 | 2.9233 | 2.3067 | 2.7517 | 3.3899 | 2.784 | 2.8167 | 2.9033 | 2.37 | 2.6609 | 2.9883
| Median |  | 2.3447 | 2.7265 | 2.78 | 2.945 | 2.72 | 2.765 | 2.73 | 2.82 | 2.925 | 2.3 | 2.755 | 3.4282 | 2.79 | 2.84 | 2.91 | 2.435 | 2.6515 | 3.005
| Std Dev. |  | 0.20494 | 0.077562 | 0.050067 | 0.046332 | 0.051284 | 0.064936 | 0.044008 | 0.033862 | 0.036697 | 0.028048 | 0.018348 | 0.080928 | 0.018864 | 0.097091 | 0.046762 | 0.48253 | 0.024503 | 0.18444
| Rel.Std.Dev. |  | 8.609% | 2.859% | 1.803% | 1.578% | 1.882% | 2.346% | 1.617% | 1.204% | 1.255% | 1.216% | 0.6668% | 2.387% | 0.6776% | 3.447% | 1.611% | 20.36% | 0.9208% | 6.172%
| PDM3 |  | -14.99% | -3.107% | -0.8452% | 4.868% | -2.69% | -1.143% | -2.809% | 0.4641% | 4.392% | -17.63% | -1.738% | 21.05% | -0.5839% | 0.5832% | 3.678% | -15.37% | -4.98% | 6.713%
| Z-Score (Absolute) |  | 3.09 | 0.47 | 0.03 | 1.3 | 0.38 | 0.03 | 0.4 | 0.32 | 1.19 | 3.68 | 0.16 | 4.88 | 0.09 | 0.35 | 1.03 | 3.18 | 0.88 | 1.7
| Analytical results for La in OREAS 290 (Certified Value 39.6 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 41.7995 | 31.2 | 37.6 | 45.4 | 36.4 | 41.2 | 40.5 | 34.9 | 37 | 38.1 | 37 | 40.66 | 33.1 | 38.5 | 43.9193 | 41 |  |  | 1
| 1 | 2 | 39.4597 | 34.4 | 36.6 | 46.5 | 38.6 | 42.7 | 42 | 27.4 | 37.6 | 40 | 39 | 40.47 | 34.5 | 41.2 | 44.4702 | 40.7 |  |  | 5
| 1 | 3 | 40.9825 | 40.3 | 37.2 | 46.6 | 36.3 | 42.5 | 41.3 | 29.9 | 37.2 | 39.8 | 37.5 | 40.65 | 33.5 | 40.4 | 44.9762 | 37.4 |  |  | 16
| 1 | 4 | 39.0498 | 36.8 | 37 | 44.1 | 41.9 | 43.3 | 41.3 | 24 | 36.6 | 39.9 | 37.3 | 41.12 | 35.2 | 42.4 | 44.7365 | 40.3 |  |  | 39.5905
| 1 | 5 | 39.0249 | 32.9 | 37.8 | 45 | 34.8 | 40.2 | 40.8 | 25.9 | 37.1 | 39.6 | 37 | 40.79 | 35.2 | 41 | 44.9898 | 41.2 |  |  | 37
| 1 | 6 | 42.7007 | 32.6 | 36.6 | 44.9 | 37.3 | 42.2 | 40.4 | 25.3 | 37.1 | 40.6 | 37.5 | 40.35 | 35.6 | 40.8 | 44.5179 | 40
| Mean |  | 40.5028 | 34.7 | 37.1333 | 45.4167 | 37.55 | 42.0167 | 41.05 | 27.9 | 37.1 | 39.6667 | 37.55 | 40.6733 | 34.5167 | 40.7167 | 44.6016 | 40.1
| Median |  | 40.2211 | 33.65 | 37.1 | 45.2 | 36.85 | 42.35 | 41.05 | 26.65 | 37.1 | 39.85 | 37.4 | 40.655 | 34.85 | 40.9 | 44.6272 | 40.5
| Std Dev. |  | 1.5573 | 3.3395 | 0.50067 | 0.97451 | 2.4712 | 1.1268 | 0.60249 | 3.9805 | 0.32249 | 0.83825 | 0.74498 | 0.26823 | 1.0147 | 1.2781 | 0.39978 | 1.3943
| Rel.Std.Dev. |  | 3.845% | 9.624% | 1.348% | 2.146% | 6.581% | 2.682% | 1.468% | 14.27% | 0.8692% | 2.113% | 1.984% | 0.6595% | 2.94% | 3.139% | 0.8963% | 3.477%
| PDM3 |  | 2.304% | -12.35% | -6.207% | 14.72% | -5.154% | 6.128% | 3.686% | -29.53% | -6.291% | 0.1923% | -5.154% | 2.735% | -12.82% | 2.844% | 12.66% | 1.287%
| Z-Score (Absolute) |  | 0.18 | 1.5 | 0.79 | 1.6 | 0.67 | 0.62 | 0.34 | 3.46 | 0.8 | 0.06 | 0.67 | 0.23 | 1.55 | 0.24 | 1.36 | 0.06
| Analytical results for Li in OREAS 290 (Certified Value 51 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
1
| 1 | 1 | 49.2 | 48.8 | 56.7 | 48.5 | 48.9 | 48.2 | 55 | 53 | 46 | 51 | 52.4767 | 50.3 | 52.3 | 52 | 48.9 | 53.4317 | 47.7 |  | 1
| 1 | 2 | 49.4 | 46.4 | 57.7 | 53.6 | 52.2 | 48.7 | 53 | 52 | 45 | 51 | 52.3585 | 49.9 | 54.4 | 53 | 50.3 | 54.57 | 50 |  | 6
| 1 | 3 | 48.7 | 51.2 | 59.8 | 49.8 | 51.2 | 47.3 | 54 | 53 | 47 | 52 | 53.1832 | 49.3 | 53.2 | 52 | 47.9 | 52.69 | 48.6 |  | 16
| 1 | 4 | 49.8 | 51.9 | 57.4 | 51.3 | 53.5 | 48.8 | 53 | 53 | 46 | 52 | 52.9599 | 50.7 | 52.3 | 53 | 53.5 | 53.86 | 50.8 |  | 51.3773
| 1 | 5 | 47 | 52.8 | 56.7 | 44.7 | 48.5 | 47 | 52 | 52 | 48 | 52 | 53.1515 | 50.8 | 52.9 | 54 | 50.9 | 53.1 | 49.2 |  | 38
| 1 | 6 | 49 | 50.8 | 57.8 | 48.4 | 52.8 | 47.6 | 54 | 54 | 49 | 51 | 53.1105 | 50.2 | 54.4 | 54 | 52.8 | 53.9753 | 50
| Mean |  | 48.85 | 50.3167 | 57.6833 | 49.3833 | 51.1833 | 47.9333 | 53.5 | 52.8333 | 46.8333 | 51.5 | 52.8734 | 50.2 | 53.25 | 53 | 50.7167 | 53.6045 | 49.3833
| Median |  | 49.1 | 51 | 57.55 | 49.15 | 51.7 | 47.9 | 53.5 | 53 | 46.5 | 51.5 | 53.0352 | 50.25 | 53.05 | 53 | 50.6 | 53.6458 | 49.6
| Std Dev. |  | 0.97929 | 2.3379 | 1.1409 | 3.0103 | 2.0702 | 0.74744 | 1.0488 | 0.75277 | 1.472 | 0.54772 | 0.36319 | 0.55136 | 0.95656 | 0.89443 | 2.1693 | 0.67166 | 1.1179
| Rel.Std.Dev. |  | 2.005% | 4.646% | 1.978% | 6.096% | 4.045% | 1.559% | 1.96% | 1.425% | 3.143% | 1.064% | 0.6869% | 1.098% | 1.796% | 1.688% | 4.277% | 1.253% | 2.264%
| PDM3 |  | -4.919% | -2.064% | 12.27% | -3.881% | -0.3776% | -6.703% | 4.132% | 2.834% | -8.844% | 0.2388% | 2.912% | -2.292% | 3.645% | 3.158% | -1.286% | 4.335% | -3.881%
| Z-Score (Absolute) |  | 0.87 | 0.32 | 2.44 | 0.67 | 0 | 1.22 | 0.87 | 0.62 | 1.63 | 0.12 | 0.63 | 0.37 | 0.77 | 0.68 | 0.17 | 0.91 | 0.67
| Analytical results for Lu in OREAS 290 (Certified Value 0.24 ppm) | [19 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.03 | 2.04 | 2.09 | 2.11 | 2.14 | 2.16 | 2.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.2 | 0.2 | 0.21 | 0.2 | 0.28 | 0.2 | 0.27 | [11 empty] | 1
| 1 | 2 | 0.2 | 0.2 | 0.22 | 0.24 | 0.3 | 0.25 | 0.27 | [11 empty] | 18
| 1 | 3 | 0.3 | 0.2 | 0.24 | 0.23 | 0.3 | 0.22 | 0.26 | [11 empty] | 16
| 1 | 4 | 0.2 | 0.2 | 0.19 | 0.22 | 0.3 | 0.23 | 0.26 | [11 empty] | 0.23619
| 1 | 5 | 0.2 | 0.2 | 0.18 | 0.25 | 0.3 | 0.25 | 0.28 | [11 empty] | 39
| 1 | 6 | 0.3 | 0.2 | 0.19 | 0.24 | 0.29 | 0.2 | 0.25
| Mean |  | 0.23333 | 0.2 | 0.205 | 0.23 | 0.295 | 0.225 | 0.265
| Median |  | 0.2 | 0.2 | 0.2 | 0.235 | 0.3 | 0.225 | 0.265
| Std Dev. |  | 0.05164 | 3.0405e-17 | 0.022583 | 0.017889 | 0.0083666 | 0.022583 | 0.010488
| Rel.Std.Dev. |  | 22.13% | 1.52e-14% | 11.02% | 7.778% | 2.836% | 10.04% | 3.958%
| PDM3 |  | -1.21% | -15.32% | -13.21% | -2.621% | 24.9% | -4.738% | 12.2%
| Z-Score (Absolute) |  | 0.09 | 0.81 | 0.67 | 0 | 1.75 | 0.13 | 0.94
| Analytical results for Mg in OREAS 290 (Certified Value 1.39 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.5079 | 1.365 | 1.32 | 1.42 | 1.31 | 1.32 | 1.26 | 1.42 | 1.51 | 1.29 | 1.74 | 1.4179 | 1.415 | 1.41 | 1.35 | 1.44 | 1.316 | 1.4 | 1
| 1 | 2 | 1.4102 | 1.408 | 1.34 | 1.44 | 1.32 | 1.36 | 1.28 | 1.35 | 1.51 | 1.29 | 1.68 | 1.4201 | 1.4196 | 1.5 | 1.39 | 1.45 | 1.33 | 1.46 | 7
| 1 | 3 | 1.5137 | 1.418 | 1.34 | 1.48 | 1.28 | 1.39 | 1.3 | 1.36 | 1.5 | 1.28 | 1.69 | 1.4311 | 1.405 | 1.5 | 1.38 | 1.44 | 1.3213 | 1.33 | 16
| 1 | 4 | 1.4222 | 1.427 | 1.36 | 1.44 | 1.42 | 1.36 | 1.29 | 1.3 | 1.47 | 1.25 | 1.66 | 1.4303 | 1.4115 | 1.44 | 1.37 | 1.44 | 1.335 | 1.29 | 1.388
| 1 | 5 | 1.4171 | 1.345 | 1.39 | 1.43 | 1.34 | 1.32 | 1.31 | 1.29 | 1.48 | 1.28 | 1.68 | 1.4354 | 1.4255 | 1.55 | 1.42 | 1.39 | 1.335 | 1.49 | 40
| 1 | 6 | 1.5409 | 1.388 | 1.37 | 1.45 | 1.29 | 1.4 | 1.28 | 1.32 | 1.46 | 1.29 | 1.7 | 1.4292 | 1.4275 | 1.48 | 1.4 | 1.37 | 1.3125 | 1.5
| Mean |  | 1.4687 | 1.3918 | 1.3533 | 1.4433 | 1.3267 | 1.3583 | 1.2867 | 1.34 | 1.4883 | 1.28 | 1.6917 | 1.4273 | 1.4174 | 1.48 | 1.385 | 1.4217 | 1.325 | 1.4117
| Median |  | 1.4651 | 1.398 | 1.35 | 1.44 | 1.315 | 1.36 | 1.285 | 1.335 | 1.49 | 1.285 | 1.685 | 1.4297 | 1.4173 | 1.49 | 1.385 | 1.44 | 1.3257 | 1.43
| Std Dev. |  | 0.058366 | 0.031971 | 0.025033 | 0.020656 | 0.050465 | 0.033714 | 0.017512 | 0.047749 | 0.02137 | 0.015492 | 0.027142 | 0.006834 | 0.0085648 | 0.049396 | 0.02429 | 0.033116 | 0.0097509 | 0.087044
| Rel.Std.Dev. |  | 3.974% | 2.297% | 1.85% | 1.431% | 3.804% | 2.482% | 1.361% | 3.563% | 1.436% | 1.21% | 1.604% | 0.4788% | 0.6043% | 3.338% | 1.754% | 2.329% | 0.7359% | 6.166%
| PDM3 |  | 5.809% | 0.2727% | -2.501% | 3.983% | -4.422% | -2.141% | -7.304% | -3.462% | 7.225% | -7.784% | 21.87% | 2.829% | 2.111% | 6.624% | -0.2196% | 2.422% | -4.544% | 1.702%
| Z-Score (Absolute) |  | 0.82 | 0.12 | 0.59 | 0.51 | 0.92 | 0.53 | 1.41 | 0.76 | 1.06 | 1.49 | 3.55 | 0.31 | 0.19 | 0.96 | 0.21 | 0.24 | 0.94 | 0.12
| Analytical results for Mn in OREAS 290 (Certified Value 0.056 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.057716 | 0.0541 | 0.0546 | 0.056 | 0.0529 | 0.0516 | 0.0541 | 0.05 | 0.0577 | 0.0575 | 0.0562 | 0.05683 | 0.0547 | 0.0544 | 0.0566 | 0.0595 | 0.060699 | 0.0554 | 1
| 1 | 2 | 0.053969 | 0.0543 | 0.0546 | 0.0565 | 0.0536 | 0.0533 | 0.0543 | 0.05 | 0.0569 | 0.0568 | 0.0556 | 0.058865 | 0.0549 | 0.0575 | 0.0587 | 0.0599 | 0.061078 | 0.0566 | 8
| 1 | 3 | 0.057711 | 0.0541 | 0.0553 | 0.0567 | 0.0517 | 0.0547 | 0.0556 | 0.05 | 0.0574 | 0.0563 | 0.0561 | 0.061278 | 0.0546 | 0.057 | 0.0581 | 0.0585 | 0.060457 | 0.0552 | 16
| 1 | 4 | 0.054828 | 0.055 | 0.056 | 0.057 | 0.054 | 0.0533 | 0.0554 | 0.05 | 0.0567 | 0.057 | 0.0562 | 0.060944 | 0.0549 | 0.0546 | 0.0577 | 0.0574 | 0.06093 | 0.0563 | 0.056358
| 1 | 5 | 0.054226 | 0.0528 | 0.057 | 0.0569 | 0.0547 | 0.0519 | 0.0555 | 0.05 | 0.0562 | 0.0577 | 0.0552 | 0.056827 | 0.0549 | 0.0587 | 0.0598 | 0.0595 | 0.060915 | 0.0571 | 41
| 1 | 6 | 0.058642 | 0.0547 | 0.0558 | 0.0568 | 0.0524 | 0.0554 | 0.0548 | 0.05 | 0.0564 | 0.0573 | 0.0556 | 0.058546 | 0.055 | 0.0562 | 0.0588 | 0.056 | 0.060687 | 0.0587
| Mean |  | 0.056182 | 0.054167 | 0.05555 | 0.05665 | 0.053217 | 0.053367 | 0.05495 | 0.05 | 0.056883 | 0.0571 | 0.055817 | 0.058882 | 0.054833 | 0.0564 | 0.058283 | 0.058467 | 0.060794 | 0.05655
| Median |  | 0.05627 | 0.0542 | 0.05555 | 0.05675 | 0.05325 | 0.0533 | 0.0551 | 0.05 | 0.0568 | 0.05715 | 0.05585 | 0.058705 | 0.0549 | 0.0566 | 0.0584 | 0.059 | 0.060807 | 0.05645
| Std Dev. |  | 0.0020641 | 0.00075807 | 0.00092033 | 0.00036194 | 0.001098 | 0.0014962 | 0.0006473 | 7.6012e-18 | 0.00057764 | 0.0005099 | 0.00041191 | 0.0019254 | 0.00015055 | 0.0016817 | 0.0010907 | 0.0015108 | 0.00022278 | 0.0012755
| Rel.Std.Dev. |  | 3.674% | 1.4% | 1.657% | 0.6389% | 2.063% | 2.804% | 1.178% | 1.52e-14% | 1.015% | 0.893% | 0.738% | 3.27% | 0.2746% | 2.982% | 1.871% | 2.584% | 0.3665% | 2.256%
| PDM3 |  | -0.3127% | -3.889% | -1.434% | 0.5176% | -5.574% | -5.308% | -2.499% | -11.28% | 0.9316% | 1.316% | -0.9611% | 4.477% | -2.706% | 0.07396% | 3.416% | 3.741% | 7.871% | 0.3401%
| Z-Score (Absolute) |  | 0.05 | 1.02 | 0.36 | 0.17 | 1.48 | 1.41 | 0.65 | 3.03 | 0.29 | 0.39 | 0.23 | 1.25 | 0.7 | 0.05 | 0.96 | 1.05 | 2.17 | 0.12
| Analytical results for Mo in OREAS 290 (Certified Value 1.14 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
2
| 1 | 1 | < 2 | 0.44 | 0.83 | 1.15 | 1.11 | 1.16 | 1.27 | 0.29 | < 1 | 1.3 | < 1 | 1.1885 | 1.2 | 1.4 | 1 | 0.11 | 2.5982 | 0.9 | 1
| 1 | 2 | < 2 | 0.44 | 0.75 | 1.2 | 1.25 | 1.24 | 1.27 | 0.33 | < 1 | 1.39 | < 1 | 1.2125 | 1.1 | 1.3 | 1 | 0.12 | 2.4954 | 0.8 | 9
| 1 | 3 | < 2 | 0.57 | 0.94 | 1.27 | 1.16 | 1.22 | 1.16 | 0.31 | < 1 | 1.3 | < 1 | 1.2069 | 1.4 | 1.2 | 1 | 0.1 | 2.5617 | 0.8 | 16
| 1 | 4 | < 2 | 0.52 | 1.01 | 1.19 | 1.22 | 1.31 | 1.22 | 0.36 | < 1 | 1.28 | 1 | 1.1928 | 1.2 | 1.2 | 1 | 0.11 | 2.5227 | 0.9 | 1.1356
| 1 | 5 | < 2 | 0.44 | 0.98 | 1.42 | 1.06 | 1.4 | 1.22 | 0.4 | < 1 | 1.28 | < 1 | 1.2013 | 1.1 | 1.4 | 1 | 0.08 | 2.5735 | 0.8 | 42
| 1 | 6 | < 2 | 0.56 | 0.86 | 1.14 | 1.48 | 1.33 | 1.23 | 0.36 | < 1 | 1.3 | 1 | 1.1873 | 1.2 | 1.2 | < 1 | 0.08 | 2.5376 | 0.9
| Mean |  |  | 0.495 | 0.895 | 1.2283 | 1.2133 | 1.2767 | 1.2283 | 0.34167 |  | 1.3083 | 1 | 1.1982 | 1.2 | 1.2833 | 1 | 0.1 | 2.5482 | 0.85
| Median |  |  | 0.48 | 0.9 | 1.195 | 1.19 | 1.275 | 1.225 | 0.345 |  | 1.3 | 1 | 1.197 | 1.2 | 1.25 | 1 | 0.105 | 2.5496 | 0.85
| Std Dev. |  |  | 0.06253 | 0.098944 | 0.10458 | 0.14801 | 0.08641 | 0.040702 | 0.039707 |  | 0.041191 | 0 | 0.010317 | 0.10954 | 0.098319 | 0 | 0.016733 | 0.037066 | 0.054772
| Rel.Std.Dev. |  |  | 12.63% | 11.06% | 8.514% | 12.2% | 6.768% | 3.314% | 11.62% |  | 3.148% | 0% | 0.861% | 9.129% | 7.661% | 0% | 16.73% | 1.455% | 6.444%
| PDM3 |  |  | -56.41% | -21.19% | 8.169% | 6.848% | 12.42% | 8.169% | -69.91% |  | 15.21% | -11.94% | 5.516% | 5.673% | 13.01% | -11.94% | -91.19% | 124.4% | -25.15%
| Z-Score (Absolute) |  | 0.34 | 2.1 | 0.71 | 0.45 | 0.4 | 0.62 | 0.45 | 2.63 | 2.08 | 0.73 | 1.5 | 0.34 | 0.35 | 0.64 | 0.63 | 3.47 | 5.03 | 0.87
| Analytical results for Na in OREAS 290 (Certified Value 0.613 wt.%) | [19 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.64103 | 0.599 | 0.586 | 0.62 | 0.58 | 0.58 | 0.59 | 0.63 | 0.65 | 0.89 | 0.59 | 0.61977 | 0.603 | 0.614 | 0.61 | 0.61 | 0.93444 | 0.424 | 1
| 1 | 2 | 0.60202 | 0.609 | 0.59 | 0.63 | 0.58 | 0.58 | 0.59 | 0.63 | 0.64 | 0.87 | 0.59 | 0.62317 | 0.6019 | 0.657 | 0.64 | 0.61 | 0.915 | 0.461 | #N/A
| 1 | 3 | 0.64481 | 0.6 | 0.592 | 0.64 | 0.56 | 0.59 | 0.6 | 0.65 | 0.63 | 0.86 | 0.6 | 0.65 | 0.5959 | 0.659 | 0.63 | 0.61 | 0.94 | 0.431 | 16
| 1 | 4 | 0.60919 | 0.61 | 0.608 | 0.63 | 0.6 | 0.59 | 0.59 | 0.63 | 0.63 | 0.9 | 0.6 | 0.65566 | 0.5939 | 0.625 | 0.63 | 0.6 | 0.93 | 0.368 | 0.61344
| 1 | 5 | 0.60501 | 0.595 | 0.617 | 0.63 | 0.59 | 0.57 | 0.6 | 0.6 | 0.64 | 0.89 | 0.6 | 0.64104 | 0.5993 | 0.677 | 0.65 | 0.58 | 0.93 | 0.483 | 43
| 1 | 6 | 0.65637 | 0.599 | 0.609 | 0.63 | 0.57 | 0.6 | 0.6 | 0.63 | 0.64 | 0.88 | 0.6 | 0.63778 | 0.6009 | 0.649 | 0.64 | 0.57 | 0.94 | 0.506
| Mean |  | 0.62641 | 0.602 | 0.60033 | 0.63 | 0.58 | 0.585 | 0.595 | 0.62833 | 0.63833 | 0.88167 | 0.59667 | 0.6379 | 0.59915 | 0.64683 | 0.63333 | 0.59667 | 0.93157 | 0.4455
| Median |  | 0.62511 | 0.5995 | 0.6 | 0.63 | 0.58 | 0.585 | 0.595 | 0.63 | 0.64 | 0.885 | 0.6 | 0.63941 | 0.6001 | 0.653 | 0.635 | 0.605 | 0.93222 | 0.446
| Std Dev. |  | 0.023662 | 0.0060663 | 0.012596 | 0.0063246 | 0.014142 | 0.010488 | 0.0054772 | 0.016021 | 0.0075277 | 0.01472 | 0.005164 | 0.014263 | 0.0035664 | 0.023327 | 0.013663 | 0.017512 | 0.0092724 | 0.04896
| Rel.Std.Dev. |  | 3.777% | 1.008% | 2.098% | 1.004% | 2.438% | 1.793% | 0.9205% | 2.55% | 1.179% | 1.67% | 0.8655% | 2.236% | 0.5952% | 3.606% | 2.157% | 2.935% | 0.9953% | 10.99%
| PDM3 |  | 2.114% | -1.865% | -2.137% | 2.699% | -5.451% | -4.636% | -3.006% | 2.428% | 4.058% | 43.72% | -2.735% | 3.987% | -2.33% | 5.443% | 3.243% | -2.735% | 51.86% | -27.38%
| Z-Score (Absolute) |  | 0.43 | 0.43 | 0.49 | 0.56 | 1.2 | 1.03 | 0.68 | 0.5 | 0.85 | 9.41 | 0.62 | 0.83 | 0.53 | 1.15 | 0.67 | 0.62 | 11.17 | 5.94
| Analytical results for Nb in OREAS 290 (Certified Value 10.2 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 9.2546 | 6.26 | 8.7 | 13.8 | 11.8 | 13.7 | 13.8 | 0.3 | 11.3 | 12.7509 | 12.08 | 11.8 | 8.6 | 0.8 | 6.3074 | 11 |  |  | 1
| 1 | 2 | 6.0521 | 7.24 | 8.69 | 13.6 | 12.4 | 14.5 | 13.8 | 3.4 | 10.6 | 12.7198 | 12.04 | 11.4 | 8 | 0.1 | 6.6236 | 10.7 |  |  | 10
| 1 | 3 | 7.9277 | 8.51 | 9.89 | 14 | 11.8 | 14.1 | 13 | 3.3 | 10.5 | 12.8493 | 12.19 | 9.8 | 8.6 | < 0.1 | 6.7642 | 10 |  |  | 16
| 1 | 4 | 6.6032 | 9.46 | 10.32 | 13.5 | 12.1 | 14.8 | 13.3 | 5.1 | 10.2 | 12.9752 | 12.34 | 9.6 | 8.7 | < 0.1 | 6.64 | 9.9 |  |  | 10.1868
| 1 | 5 | 6.6487 | 6.23 | 10.01 | 13.7 | 11 | 13.7 | 13 | 3.5 | 11.2 | 13.0193 | 12.44 | 9.6 | 8.4 | < 0.1 | 6.6 | 10.6 |  |  | 44
| 1 | 6 | 5.5936 | 8.26 | 9.71 | 13.4 | 11.4 | 14.5 | 12.5 | 4.9 | 10.5 | 12.6588 | 11.91 | 9.8 | 9.1 | < 0.1 | 6.5 | 10.7
| Mean |  | 7.0133 | 7.66 | 9.5533 | 13.6667 | 11.75 | 14.2167 | 13.2333 | 3.4167 | 10.7167 | 12.8289 | 12.1667 | 10.3333 | 8.5667 | 0.45 | 6.5725 | 10.4833
| Median |  | 6.6259 | 7.75 | 9.8 | 13.65 | 11.8 | 14.3 | 13.15 | 3.45 | 10.55 | 12.8001 | 12.135 | 9.8 | 8.6 | 0.45 | 6.6118 | 10.65
| Std Dev. |  | 1.3491 | 1.3041 | 0.69394 | 0.21602 | 0.49699 | 0.45789 | 0.50859 | 1.7186 | 0.43551 | 0.14491 | 0.19735 | 0.99331 | 0.36148 | 0.49497 | 0.15506 | 0.43551
| Rel.Std.Dev. |  | 19.24% | 17.03% | 7.264% | 1.581% | 4.23% | 3.221% | 3.843% | 50.3% | 4.064% | 1.13% | 1.622% | 9.613% | 4.22% | 110% | 2.359% | 4.154%
| PDM3 |  | -31.15% | -24.8% | -6.218% | 34.16% | 15.35% | 39.56% | 29.91% | -66.46% | 5.202% | 25.94% | 19.44% | 1.439% | -15.9% | -95.58% | -35.48% | 2.911%
| Z-Score (Absolute) |  | 0.89 | 0.72 | 0.22 | 0.85 | 0.35 | 0.99 | 0.74 | 1.82 | 0.08 | 0.63 | 0.46 | 0.02 | 0.48 | 2.67 | 1 | 0.02
| Analytical results for Nd in OREAS 290 (Certified Value 34.2 ppm) | [19 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 1
2
| 1 | 1 | 32.4 | 32.4 | 35.77 | 30.3 | 36.6 | [13 empty] | 1
| 1 | 2 | 35 | 31.6 | 34.88 | 31.6 | 36.7 | [13 empty] | 19
| 1 | 3 | 38.5 | 32.5 | 35.81 | 30.6 | 33.1 | [13 empty] | 16
| 1 | 4 | 37 | 32.3 | 35.53 | 32.9 | 35.5 | [13 empty] | 34.2123
| 1 | 5 | 32.1 | 33.1 | 36.07 | 32.1 | 36.3 | [13 empty] | 45
| 1 | 6 | 36.8 | 31.3 | 35.41 | 33.2 | 35.9
| Mean |  | 35.3 | 32.2 | 35.5783 | 31.7833 | 35.6833
| Median |  | 35.9 | 32.35 | 35.65 | 31.85 | 36.1
| Std Dev. |  | 2.6123 | 0.65115 | 0.41243 | 1.1822 | 1.3423
| Rel.Std.Dev. |  | 7.4% | 2.022% | 1.159% | 3.72% | 3.762%
| PDM3 |  | 3.179% | -5.882% | 3.993% | -7.1% | 4.3%
| Z-Score (Absolute) |  | 0 | 5.45 | 0.49 | 6.19 | 0.67
| Analytical results for Ni in OREAS 290 (Certified Value 73 ppm) | [19 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 80.9853 | 72.7 | 71 | 75.9 | 69.4 | 70.9 | 70.6 | 69.8 | 73 | 71 | 77 | 74.218 | 75.7 | 74 | 67 | 85.3 | 69.9864 | 69 | 1
| 1 | 2 | 75.8637 | 73.2 | 69.5 | 76 | 72.1 | 74 | 70.7 | 68.8 | 74 | 69.7 | 75 | 74.6669 | 74.1 | 77 | 78 | 73.8 | 71.1854 | 69.1 | 11
| 1 | 3 | 81.5188 | 72.1 | 71.4 | 77.9 | 68.4 | 73.3 | 71.7 | 69 | 73 | 70.7 | 76 | 76.1005 | 71.9 | 78 | 70 | 70.7 | 70.8035 | 65.3 | 16
| 1 | 4 | 77.3575 | 72.8 | 72 | 74.2 | 71 | 73.6 | 72.5 | 68.6 | 73 | 70.7 | 74 | 76.8125 | 72.6 | 75 | 69 | 76.3 | 70.2471 | 65.7 | 72.591
| 1 | 5 | 76.5039 | 74.1 | 73.7 | 76.5 | 64.9 | 67.5 | 70.9 | 68.2 | 73 | 71.9 | 75 | 76.7606 | 72.9 | 80 | 76 | 71.7 | 69.9743 | 64.6 | 46
| 1 | 6 | 83.0126 | 72.1 | 73.4 | 76.1 | 66.8 | 71.4 | 70.1 | 70.6 | 72 | 69.8 | 71 | 75.8396 | 72.7 | 74 | 71 | 75.8 | 71.0952 | 66
| Mean |  | 79.207 | 72.8333 | 71.8333 | 76.1 | 68.7667 | 71.7833 | 71.0833 | 69.1667 | 73 | 70.6333 | 74.6667 | 75.733 | 73.3167 | 76.3333 | 71.8333 | 75.6 | 70.5486 | 66.6167
| Median |  | 79.1714 | 72.75 | 71.7 | 76.05 | 68.9 | 72.35 | 70.8 | 68.9 | 73 | 70.7 | 75 | 75.9701 | 72.8 | 76 | 70.5 | 74.8 | 70.5253 | 65.85
| Std Dev. |  | 2.9965 | 0.75277 | 1.568 | 1.1883 | 2.6628 | 2.4408 | 0.86814 | 0.88015 | 0.63246 | 0.81404 | 2.0656 | 1.077 | 1.3688 | 2.4221 | 4.2622 | 5.236 | 0.54883 | 1.9426
| Rel.Std.Dev. |  | 3.783% | 1.034% | 2.183% | 1.561% | 3.872% | 3.4% | 1.221% | 1.273% | 0.8664% | 1.152% | 2.766% | 1.422% | 1.867% | 3.173% | 5.934% | 6.926% | 0.7779% | 2.916%
| PDM3 |  | 9.114% | 0.3338% | -1.044% | 4.834% | -5.268% | -1.113% | -2.077% | -4.717% | 0.5634% | -2.697% | 2.859% | 4.328% | 0.9996% | 5.155% | -1.044% | 4.145% | -2.814% | -8.23%
| Z-Score (Absolute) |  | 2.25 | 0.16 | 0.16 | 1.23 | 1.17 | 0.18 | 0.41 | 1.04 | 0.22 | 0.56 | 0.76 | 1.11 | 0.32 | 1.31 | 0.16 | 1.07 | 0.58 | 1.87
| Analytical results for P in OREAS 290 (Certified Value 0.059 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.061334 | 0.0501 | 0.055 | 0.061 | 0.058 | 0.056 | 0.059 | 0.0496 | 0.06 | 0.0538 | 0.05 | 0.059141 | 0.0609 | 0.059 | 0.06 | 0.05 | 0.062856 | 0.0576 | 1
| 1 | 2 | 0.056676 | 0.0528 | 0.052 | 0.062 | 0.058 | 0.057 | 0.058 | 0.0495 | 0.06 | 0.0529 | 0.05 | 0.057756 | 0.0599 | 0.063 | 0.06 | 0.05 | 0.062191 | 0.0623 | 12
| 1 | 3 | 0.060579 | 0.0562 | 0.056 | 0.064 | 0.056 | 0.059 | 0.059 | 0.0505 | 0.06 | 0.0539 | 0.06 | 0.058424 | 0.06 | 0.061 | 0.06 | 0.049 | 0.062646 | 0.0603 | 16
| 1 | 4 | 0.056565 | 0.0562 | 0.058 | 0.062 | 0.06 | 0.058 | 0.058 | 0.0497 | 0.06 | 0.0535 | 0.06 | 0.058358 | 0.0603 | 0.058 | 0.06 | 0.048 | 0.062522 | 0.0507 | 0.05857
| 1 | 5 | 0.056513 | 0.051 | 0.058 | 0.061 | 0.06 | 0.056 | 0.058 | 0.0493 | 0.06 | 0.0551 | 0.06 | 0.058982 | 0.061 | 0.063 | 0.06 | 0.047 | 0.062512 | 0.0599 | 47
| 1 | 6 | 0.06126 | 0.0538 | 0.058 | 0.062 | 0.058 | 0.06 | 0.057 | 0.0514 | 0.06 | 0.0567 | 0.06 | 0.058189 | 0.0607 | 0.06 | 0.06 | 0.049 | 0.062397 | 0.0598
| Mean |  | 0.058821 | 0.05335 | 0.056167 | 0.062 | 0.058333 | 0.057667 | 0.058167 | 0.05 | 0.06 | 0.054317 | 0.056667 | 0.058475 | 0.060467 | 0.060667 | 0.06 | 0.048833 | 0.062521 | 0.058433
| Median |  | 0.058627 | 0.0533 | 0.057 | 0.062 | 0.058 | 0.0575 | 0.058 | 0.04965 | 0.06 | 0.05385 | 0.06 | 0.058391 | 0.0605 | 0.0605 | 0.06 | 0.049 | 0.062517 | 0.05985
| Std Dev. |  | 0.0024646 | 0.0025626 | 0.0024014 | 0.0010954 | 0.0015055 | 0.001633 | 0.00075277 | 0.0008 | 0 | 0.0013717 | 0.005164 | 0.00051302 | 0.00046762 | 0.0020656 | 0 | 0.001169 | 0.00022456 | 0.0040732
| Rel.Std.Dev. |  | 4.19% | 4.803% | 4.275% | 1.767% | 2.581% | 2.832% | 1.294% | 1.6% | 0% | 2.525% | 9.113% | 0.8773% | 0.7733% | 3.405% | 0% | 2.394% | 0.3592% | 6.971%
| PDM3 |  | 0.4288% | -8.912% | -4.103% | 5.856% | -0.4041% | -1.542% | -0.6887% | -14.63% | 2.441% | -7.262% | -3.25% | -0.162% | 3.238% | 3.58% | 2.441% | -16.62% | 6.745% | -0.2334%
| Z-Score (Absolute) |  | 0.16 | 1.79 | 0.79 | 1.28 | 0.02 | 0.25 | 0.08 | 2.98 | 0.57 | 1.44 | 0.61 | 0.03 | 0.74 | 0.81 | 0.57 | 3.39 | 1.47 | 0.02
| Analytical results for Pb in OREAS 290 (Certified Value 22.9 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 19.71 | 22.48 | 24.2 | 22.4 | 21.9 | 21.7 | 23.6 | 23 | 23.1 | 20 | 29.1467 | 24.6 | 22 | < 50 | 24.4 | 22.7711 | 24.8 |  | 1
| 1 | 2 | 21.89 | 22.17 | 23.9 | 22.3 | 22.1 | 22.9 | 22.8 | 23 | 22.4 | 24 | 27.8765 | 24.5 | 23 | < 50 | 24.8 | 23.1608 | 25.6 |  | 13
| 1 | 3 | 20.9 | 22.76 | 24.5 | 20.8 | 23.5 | 22.7 | 23.7 | 23 | 23 | 21 | 27.4885 | 25.3 | 22 | < 50 | 24.8 | 23.024 | 21.7 |  | 16
| 1 | 4 | 20.86 | 22.2 | 23.5 | 21.8 | 22.2 | 23.3 | 22.6 | 24 | 22.2 | 20 | 28.8859 | 24.7 | 22 | < 50 | 25.5 | 22.9866 | 24.3 |  | 22.8867
| 1 | 5 | 21.12 | 23.11 | 23.8 | 20.4 | 20.8 | 22.7 | 22.8 | 22 | 22.6 | 21 | 26.2476 | 24.5 | 23 | < 50 | 24.8 | 23.233 | 25.2 |  | 48
| 1 | 6 | 22.91 | 21.69 | 23.7 | 21.4 | 22 | 22.3 | 23.4 | 23 | 23.9 | 20 | 25.6414 | 25.5 | 23 | < 50 | 25.1 | 22.864 | 24.4
| Mean |  | 21.2317 | 22.4017 | 23.9333 | 21.5167 | 22.0833 | 22.6 | 23.15 | 23 | 22.8667 | 21 | 27.5478 | 24.85 | 22.5 |  | 24.9 | 23.0066 | 24.3333
| Median |  | 21.01 | 22.34 | 23.85 | 21.6 | 22.05 | 22.7 | 23.1 | 23 | 22.8 | 20.5 | 27.6825 | 24.65 | 22.5 |  | 24.8 | 23.0053 | 24.6
| Std Dev. |  | 1.0795 | 0.49741 | 0.36148 | 0.80602 | 0.8612 | 0.54772 | 0.47223 | 0.63246 | 0.6121 | 1.5492 | 1.3988 | 0.43704 | 0.54772 |  | 0.36878 | 0.17408 | 1.3794
| Rel.Std.Dev. |  | 5.084% | 2.22% | 1.51% | 3.746% | 3.9% | 2.424% | 2.04% | 2.75% | 2.677% | 7.377% | 5.078% | 1.759% | 2.434% |  | 1.481% | 0.7567% | 5.669%
| PDM3 |  | -7.231% | -2.119% | 4.573% | -5.986% | -3.51% | -1.253% | 1.151% | 0.4952% | -0.08737% | -8.243% | 20.37% | 8.579% | -1.689% |  | 8.797% | 0.524% | 6.321%
| Z-Score (Absolute) |  | 1.28 | 0.43 | 0.67 | 1.07 | 0.66 | 0.29 | 0.11 | 0 | 0.1 | 1.44 | 3.29 | 1.34 | 0.36 | 1.44 | 1.37 | 0 | 0.96
| Analytical results for Pr in OREAS 290 (Certified Value 9.13 ppm) | [19 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 7.9 | 9 | 9.56 | 8.37 | 9.82 | [13 empty] | 1
| 1 | 2 | 9.2 | 8.5 | 9.5 | 8.6 | 9.76 | [13 empty] | 21
| 1 | 3 | 10.2 | 8.9 | 9.62 | 8.07 | 8.83 | [13 empty] | 16
| 1 | 4 | 9.4 | 8.9 | 9.6 | 8.5 | 9.55 | [13 empty] | 9.1335
| 1 | 5 | 8.5 | 9.2 | 9.53 | 8.49 | 9.87 | [13 empty] | 49
| 1 | 6 | 9.2 | 8.7 | 9.59 | 8.62 | 9.63
| Mean |  | 9.0667 | 8.8667 | 9.5667 | 8.4417 | 9.5767
| Median |  | 9.2 | 8.9 | 9.575 | 8.495 | 9.695
| Std Dev. |  | 0.78909 | 0.24221 | 0.045461 | 0.20292 | 0.38469
| Rel.Std.Dev. |  | 8.703% | 2.732% | 0.4752% | 2.404% | 4.017%
| PDM3 |  | -0.7321% | -2.922% | 4.742% | -7.575% | 4.852%
| Z-Score (Absolute) |  | 0 | 0.27 | 0.67 | 0.84 | 0.69
| Analytical results for Rb in OREAS 290 (Certified Value 161 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 144.1 | 159.3 | 180 | 151 | 164.5 | 146.5 | 113.5 | 169.9 | 165.1792 | 169.28 | 172 | 165.1 | 149 | 168 |  |  |  |  | 1
| 1 | 2 | 137.5 | 155.7 | 180 | 155.5 | 174 | 148 | 134.2 | 161.9 | 165.558 | 165.26 | 171 | 165.8 | 149 | 166 |  |  |  |  | 44
| 1 | 3 | 148 | 158 | 181.5 | 146.5 | 174.5 | 151 | 140.2 | 164.1 | 167.2082 | 161.9 | 168 | 165.8 | 149 | 152 |  |  |  |  | 16
| 1 | 4 | 162.3 | 160.9 | 175.5 | 155.5 | 178 | 151.5 | 122.2 | 164.5 | 166.2418 | 162.88 | 172 | 163.6 | 141 | 159 |  |  |  |  | 161.2573
| 1 | 5 | 162.3 | 163.8 | 177 | 141 | 167.5 | 151.5 | 131.3 | 164.9 | 167.3641 | 167.33 | 168 | 163.7 | 124 | 159 |  |  |  |  | 50
| 1 | 6 | 148.4 | 160.5 | 177.5 | 142 | 174 | 146.5 | 133.3 | 169.9 | 165.1057 | 163.16 | 173 | 165.7 | 129 | 167
| Mean |  | 150.4333 | 159.7 | 178.5833 | 148.5833 | 172.0833 | 149.1667 | 129.1167 | 165.8667 | 166.1095 | 164.9683 | 170.6667 | 164.95 | 140.1667 | 161.8333
| Median |  | 148.2 | 159.9 | 178.75 | 148.75 | 174 | 149.5 | 132.3 | 164.7 | 165.8999 | 164.21 | 171.5 | 165.4 | 145 | 162.5
| Std Dev. |  | 9.9907 | 2.7546 | 2.2675 | 6.4297 | 5.0341 | 2.4427 | 9.6194 | 3.2922 | 0.99758 | 2.874 | 2.1602 | 1.0407 | 11.143 | 6.2423
| Rel.Std.Dev. |  | 6.641% | 1.725% | 1.27% | 4.327% | 2.925% | 1.638% | 7.45% | 1.985% | 0.6006% | 1.742% | 1.266% | 0.6309% | 7.95% | 3.857%
| PDM3 |  | -6.712% | -0.9657% | 10.74% | -7.859% | 6.714% | -7.498% | -19.93% | 2.858% | 3.009% | 2.301% | 5.835% | 2.29% | -13.08% | 0.3572%
| Z-Score (Absolute) |  | 1.09 | 0.31 | 1.28 | 1.25 | 0.73 | 1.2 | 2.9 | 0.21 | 0.23 | 0.13 | 0.61 | 0.13 | 1.96 | 0.13
| Analytical results for Re in OREAS 290 (Certified Value < 0.002 ppm) | [19 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.05 | 0.002 | < 0.05 | < 0.02 | < 0.002 | [7 empty] | 1
| 1 | 2 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.05 | < 0.002 | < 0.05 | < 0.02 | < 0.002 | [7 empty] | 45
| 1 | 3 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.05 | < 0.002 | < 0.05 | < 0.02 | < 0.002 | [7 empty] | 16
| 1 | 4 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.05 | < 0.002 | < 0.05 | < 0.02 | < 0.002 | [7 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.05 | < 0.002 | < 0.05 | < 0.02 | < 0.002 | [7 empty] | 51
| 1 | 6 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.05 | < 0.002 | < 0.05 | < 0.02 | < 0.002
| Mean |  |  |  |  | 0.002 |  |  |  | 0.002
| Median |  |  |  |  | 0.002 |  |  |  | 0.002
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 290 (Certified Value 0.528 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.53432 | 0.5016 | 0.49 | 0.53 | 0.53 | 0.52 | 0.51 | 0.53 | 0.57 | 0.51 | 0.57 | 0.52137 | 0.5236 | 0.514 | 0.52 | 0.53 | 0.539 | 0.435 | 1
| 1 | 2 | 0.50258 | 0.5099 | 0.49 | 0.54 | 0.53 | 0.52 | 0.51 | 0.5 | 0.57 | 0.51 | 0.56 | 0.519 | 0.5257 | 0.55 | 0.54 | 0.54 | 0.55517 | 0.428 | 30
| 1 | 3 | 0.53443 | 0.5088 | 0.49 | 0.55 | 0.51 | 0.54 | 0.53 | 0.51 | 0.57 | 0.5 | 0.58 | 0.53822 | 0.519 | 0.556 | 0.52 | 0.51 | 0.56056 | 0.407 | 16
| 1 | 4 | 0.50668 | 0.514 | 0.49 | 0.54 | 0.53 | 0.54 | 0.53 | 0.5 | 0.57 | 0.51 | 0.56 | 0.5298 | 0.5189 | 0.523 | 0.52 | 0.52 | 0.56154 | 0.403 | 0.52785
| 1 | 5 | 0.50475 | 0.4944 | 0.49 | 0.54 | 0.53 | 0.51 | 0.52 | 0.49 | 0.57 | 0.51 | 0.57 | 0.53859 | 0.5278 | 0.567 | 0.53 | 0.53 | 0.55468 | 0.469 | 52
| 1 | 6 | 0.5414 | 0.5068 | 0.49 | 0.54 | 0.52 | 0.55 | 0.52 | 0.5 | 0.56 | 0.51 | 0.56 | 0.53138 | 0.5272 | 0.535 | 0.52 | 0.53 | 0.54439 | 0.463
| Mean |  | 0.52069 | 0.50592 | 0.49 | 0.54 | 0.525 | 0.53 | 0.52 | 0.505 | 0.56833 | 0.50833 | 0.56667 | 0.52973 | 0.5237 | 0.54083 | 0.525 | 0.52667 | 0.55256 | 0.43417
| Median |  | 0.5205 | 0.5078 | 0.49 | 0.54 | 0.53 | 0.53 | 0.52 | 0.5 | 0.57 | 0.51 | 0.565 | 0.53059 | 0.52465 | 0.5425 | 0.52 | 0.53 | 0.55493 | 0.4315
| Std Dev. |  | 0.017786 | 0.0069508 | 6.0809e-17 | 0.0063246 | 0.0083666 | 0.015492 | 0.0089443 | 0.013784 | 0.0040825 | 0.0040825 | 0.008165 | 0.0082286 | 0.0039547 | 0.020351 | 0.0083666 | 0.010328 | 0.0090174 | 0.027542
| Rel.Std.Dev. |  | 3.416% | 1.374% | 1.241e-14% | 1.171% | 1.594% | 2.923% | 1.72% | 2.73% | 0.7183% | 0.8031% | 1.441% | 1.553% | 0.7552% | 3.763% | 1.594% | 1.961% | 1.632% | 6.344%
| PDM3 |  | -1.355% | -4.155% | -7.17% | 2.302% | -0.5397% | 0.4076% | -1.487% | -4.329% | 7.67% | -3.697% | 7.354% | 0.3559% | -0.786% | 2.46% | -0.5397% | -0.2239% | 4.681% | -17.75%
| Z-Score (Absolute) |  | 0.19 | 0.83 | 1.53 | 0.66 | 0 | 0.22 | 0.22 | 0.87 | 1.9 | 0.73 | 1.82 | 0.21 | 0.06 | 0.69 | 0 | 0.07 | 1.21 | 3.97
| Analytical results for Sb in OREAS 290 (Certified Value 0.84 wt.%) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 3
3
| 1 | 1 | 0.7952 | 0.35399 | 0.86 | 0.868 | 0.812 | 0.806 | 0.847 | 0.8127 | 0.8444 | 0.86556 | 0.773 | 0.893 | 0.85047 | 0.0469 | 0.7214 | 0.835 |  |  | 1
| 1 | 2 | 0.8064 | > 0.4 | 0.84 | 0.881 | 0.819 | 0.832 | 0.86 | 0.8089 | 0.8542 | 0.85694 | 0.8107 | 0.924 | 0.84492 | 0.0404 | 0.73189 | 0.845 |  |  | 26
| 1 | 3 | 0.9268 | > 0.4 | 0.81 | 0.905 | 0.792 | 0.849 | 0.879 | 0.8291 | 0.8361 | 0.85931 | 0.8143 | 0.837 | 0.84338 | 0.0265 | 0.73806 | 0.821 |  |  | 16
| 1 | 4 | 0.8344 | > 0.4 | 0.76 | 0.884 | 0.833 | 0.825 | 0.869 | 0.8126 | 0.8444 | 0.87103 | 0.818 | 0.882 | 0.84692 | 0.0117 | 0.73364 | 0.838 |  |  | 0.83979
| 1 | 5 | 0.826 | 0.31833 | 0.85 | 0.876 | 0.827 | 0.799 | 0.87 | 0.8374 | 0.8334 | 0.87234 | 0.8228 | 0.814 | 0.83862 | 0.0232 | 0.72214 | 0.852 |  |  | 53
| 1 | 6 | 0.8092 | > 0.4 | 0.83 | 0.89 | 0.796 | 0.852 | 0.86 | 0.8199 | 0.8348 | 0.86195 | 0.8168 | 0.843 | 0.84757 | 0.0109 | 0.73344 | 0.827
| Mean |  | 0.833 | 0.33616 | 0.825 | 0.884 | 0.81317 | 0.82717 | 0.86417 | 0.8201 | 0.84122 | 0.86452 | 0.80927 | 0.8655 | 0.84531 | 0.0266 | 0.7301 | 0.83633
| Median |  | 0.8176 | 0.33616 | 0.835 | 0.8825 | 0.8155 | 0.8285 | 0.8645 | 0.8163 | 0.84025 | 0.86375 | 0.81555 | 0.8625 | 0.84592 | 0.02485 | 0.73267 | 0.8365
| Std Dev. |  | 0.048067 | 0.02521 | 0.036194 | 0.012696 | 0.01651 | 0.021739 | 0.011017 | 0.011107 | 0.007966 | 0.006258 | 0.018214 | 0.04104 | 0.0040733 | 0.014715 | 0.0067725 | 0.011378
| Rel.Std.Dev. |  | 5.77% | 7.499% | 4.387% | 1.436% | 2.03% | 2.628% | 1.275% | 1.354% | 0.947% | 0.7239% | 2.251% | 4.742% | 0.4819% | 55.32% | 0.9276% | 1.361%
| PDM3 |  | -0.8083% | -59.97% | -1.761% | 5.265% | -3.17% | -1.503% | 2.903% | -2.344% | 0.1701% | 2.945% | -3.634% | 3.062% | 0.658% | -96.83% | -13.06% | -0.4114%
| Z-Score (Absolute) |  | 0.1 | 17.65 | 0.18 | 1.93 | 0.6 | 0.1 | 1.22 | 0.36 | 0.4 | 1.23 | 0.74 | 1.27 | 0.54 | 28.72 | 3.57 | 0.22
| Analytical results for Sc in OREAS 290 (Certified Value 13.2 ppm) | [19 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 13.97 | 11.2 | 12.3 | 15 | 12.8 | 13.1 | 12 | 9.6 | 13.4 | 14.5 | 13.1 | 12.9558 | 14 | 13 | < 20 | 13.9 | 13.622 | 13 | 1
| 1 | 2 | 12.98 | 11.6 | 12.1 | 14.9 | 13 | 13.4 | 12.2 | 7.8 | 13.5 | 13.7 | 13 | 12.638 | 13.8 | 14 | < 20 | 12.9 | 13.9325 | 14 | 47
| 1 | 3 | 13.42 | 13.4 | 12.4 | 15.3 | 12.2 | 13.8 | 12.4 | 8.1 | 13.4 | 13.1 | 13.1 | 12.7208 | 13.4 | 14 | < 20 | 13 | 13.608 | 12 | 16
| 1 | 4 | 12.87 | 12.5 | 12.5 | 14.8 | 13.2 | 14.2 | 12.6 | 6.6 | 13.2 | 13.3 | 13.4 | 12.6359 | 13.2 | 13 | < 20 | 14.3 | 14.322 | 12 | 13.2056
| 1 | 5 | 12.21 | 10.9 | 12.7 | 14.5 | 11.4 | 13 | 12.4 | 6.4 | 13.4 | 13.4 | 13.1 | 12.9549 | 13.8 | 14 | < 20 | 14.9 | 13.6648 | 14 | 54
| 1 | 6 | 13.75 | 12 | 13 | 15 | 12 | 13.8 | 12.3 | 6.2 | 13.3 | 13.6 | 13.2 | 12.8003 | 13.2 | 13 | < 20 | 14.2 | 13.6455 | 15
| Mean |  | 13.2 | 11.9333 | 12.5 | 14.9167 | 12.4333 | 13.55 | 12.3167 | 7.45 | 13.3667 | 13.6 | 13.15 | 12.7843 | 13.5667 | 13.5 |  | 13.8667 | 13.7991 | 13.3333
| Median |  | 13.2 | 11.8 | 12.45 | 14.95 | 12.5 | 13.6 | 12.35 | 7.2 | 13.4 | 13.5 | 13.1 | 12.7605 | 13.6 | 13.5 |  | 14.05 | 13.6552 | 13.5
| Std Dev. |  | 0.64517 | 0.91579 | 0.31623 | 0.26394 | 0.68605 | 0.46368 | 0.20412 | 1.3081 | 0.10328 | 0.4899 | 0.13784 | 0.14575 | 0.34448 | 0.54772 |  | 0.78145 | 0.2831 | 1.2111
| Rel.Std.Dev. |  | 4.888% | 7.674% | 2.53% | 1.769% | 5.518% | 3.422% | 1.657% | 17.56% | 0.7727% | 3.602% | 1.048% | 1.14% | 2.539% | 4.057% |  | 5.635% | 2.052% | 9.083%
| PDM3 |  | -0.04265% | -9.635% | -5.343% | 12.96% | -5.848% | 2.608% | -6.732% | -43.58% | 1.219% | 2.986% | -0.4213% | -3.191% | 2.734% | 2.229% |  | 5.006% | 4.494% | 0.967%
| Z-Score (Absolute) |  | 0.03 | 1.29 | 0.7 | 1.81 | 0.77 | 0.39 | 0.89 | 5.94 | 0.2 | 0.44 | 0.03 | 0.41 | 0.41 | NA | 3.3 | 0.72 | 0.65 | NA
Results from laboratories 2.15 and 2.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 290 (Certified Value 2.64 ppm) | [19 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 12.8 | 2.1 | 3 | 4 | 2 | 3 | < 2 | < 2 | 2.1823 | 2.3 | 2 | < 5 | 2 | < 5 | 4.14 |  |  |  | 1
| 1 | 2 | 9 | 2.2 | 3 | 4 | 3 | 3 | < 2 | 2 | 2.331 | 2.7 | 2 | < 5 | 2 | < 5 | 4.62 |  |  |  | 48
| 1 | 3 | 1.8 | 2.3 | 3 | 3 | 3 | 3 | < 2 | 2 | 2.1718 | 2.9 | 2 | < 5 | 1 | < 5 | 3.78 |  |  |  | 16
| 1 | 4 | 9.9 | 2.2 | 3 | 3 | 3 | 3 | < 2 | 2 | 2.0951 | 3 | 2 | < 5 | 1 | < 5 | 4.23 |  |  |  | 2.6423
| 1 | 5 | 13 | 2.4 | 3 | 3 | 2 | 3 | < 2 | 3 | < 2 | 2.5 | 2 | < 5 | 1 | < 5 | 3.66 |  |  |  | 55
| 1 | 6 | 4.8 | 2.4 | 3 | 3 | 3 | 3 | < 2 | 3 | < 2 | 2.7 | 2 | < 5 | 1 | < 5 | 4.69
| Mean |  | 8.55 | 2.2667 | 3 | 3.3333 | 2.6667 | 3 |  | 2.4 | 2.195 | 2.6833 | 2 |  | 1.3333 |  | 4.1867
| Median |  | 9.45 | 2.25 | 3 | 3 | 3 | 3 |  | 2 | 2.1771 | 2.7 | 2 |  | 1 |  | 4.185
| Std Dev. |  | 4.4613 | 0.12111 | 0 | 0.5164 | 0.5164 | 0 |  | 0.54772 | 0.098594 | 0.25626 | 0 |  | 0.5164 |  | 0.42141
| Rel.Std.Dev. |  | 52.18% | 5.343% | 0% | 15.49% | 19.36% | 0% |  | 22.82% | 4.492% | 9.55% | 0% |  | 38.73% |  | 10.07%
| PDM3 |  | 223.6% | -14.22% | 13.54% | 26.15% | 0.9231% | 13.54% |  | -9.169% | -16.93% | 1.554% | -24.31% |  | -49.54% |  | 58.45%
| Z-Score (Absolute) |  | 8.16 | 0.31 | 0.67 | 1.12 | 0.22 | 0.67 | 2.02 | 0.45 | 0.95 | 0.25 | 0.67 | 0 | 1.57 | 0 | 2.27
| Analytical results for Sm in OREAS 290 (Certified Value 6.42 ppm) | [19 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 5.8 | 5.9 | 6.8 | 6 | 6.82 | [13 empty] | 1
| 1 | 2 | 6.3 | 5.8 | 6.96 | 5.7 | 6.82 | [13 empty] | 18
| 1 | 3 | 6.7 | 5.9 | 6.85 | 6 | 6.1 | [13 empty] | 16
| 1 | 4 | 6.5 | 5.9 | 6.93 | 6.5 | 6.81 | [13 empty] | 6.4231
| 1 | 5 | 6.2 | 6.1 | 6.83 | 6.2 | 6.88 | [13 empty] | 56
| 1 | 6 | 6.9 | 5.7 | 6.8 | 6.6 | 6.69
| Mean |  | 6.4 | 5.8833 | 6.8617 | 6.1667 | 6.6867
| Median |  | 6.4 | 5.9 | 6.84 | 6.1 | 6.815
| Std Dev. |  | 0.38987 | 0.13292 | 0.067946 | 0.33862 | 0.29405
| Rel.Std.Dev. |  | 6.092% | 2.259% | 0.9902% | 5.491% | 4.398%
| PDM3 |  | -0.3602% | -8.404% | 6.827% | -3.993% | 4.103%
| Z-Score (Absolute) |  | 0 | 1.22 | 1.09 | 0.55 | 0.67
| Analytical results for Sn in OREAS 290 (Certified Value 3.56 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 3.7145 | 2.1 | 2.9 | 3.9 | 3.7 | 3.5 | 3.8 | 1.5 | < 10 | 3.7 | < 10 | 3.5659 | 3.8 | 3.3 | 3 | 1.3 | < 10 | 3.8 | 1
| 1 | 2 | 2.9325 | 2.2 | 2.9 | 3.9 | 3.8 | 3.7 | 3.9 | 1.8 | < 10 | 3.8 | < 10 | 3.5499 | 3.6 | 3.5 | 4 | 1.2 | < 10 | 3.8 | 6
| 1 | 3 | 3.2545 | 2.6 | 3.1 | 4 | 3.6 | 3.7 | 3.9 | 1.8 | < 10 | 3.8 | < 10 | 3.6034 | 3.7 | 3.3 | 3 | 1.1 | < 10 | 3.4 | 16
| 1 | 4 | 2.9555 | 2.6 | 3.3 | 3.9 | 3.6 | 3.8 | 4 | 2.1 | < 10 | 3.8 | < 10 | 3.6327 | 3.5 | 2.8 | 4 | 0.9 | < 10 | 3.7 | 3.5648
| 1 | 5 | 2.99 | 2 | 3.1 | 3.8 | 3.4 | 3.6 | 3.9 | 2 | < 10 | 3.9 | < 10 | 3.5994 | 3.7 | 3.1 | 3 | 0.9 | < 10 | 3.8 | 57
| 1 | 6 | 3.059 | 2.6 | 3.1 | 4 | 3.6 | 3.7 | 3.8 | 2.2 | < 10 | 3.9 | < 10 | 3.518 | 3.7 | 3.1 | 4 | 1 | < 10 | 3.8
| Mean |  | 3.151 | 2.35 | 3.0667 | 3.9167 | 3.6167 | 3.6667 | 3.8833 | 1.9 |  | 3.8167 |  | 3.5782 | 3.6667 | 3.1833 | 3.5 | 1.0667 |  | 3.7167
| Median |  | 3.0245 | 2.4 | 3.1 | 3.9 | 3.6 | 3.7 | 3.9 | 1.9 |  | 3.8 |  | 3.5826 | 3.7 | 3.2 | 3.5 | 1.05 |  | 3.8
| Std Dev. |  | 0.29953 | 0.28107 | 0.15055 | 0.075277 | 0.13292 | 0.10328 | 0.075277 | 0.25298 |  | 0.075277 |  | 0.041555 | 0.10328 | 0.24014 | 0.54772 | 0.1633 |  | 0.16021
| Rel.Std.Dev. |  | 9.506% | 11.96% | 4.909% | 1.922% | 3.675% | 2.817% | 1.938% | 13.31% |  | 1.972% |  | 1.161% | 2.817% | 7.544% | 15.65% | 15.31% |  | 4.311%
| PDM3 |  | -11.61% | -34.08% | -13.97% | 9.87% | 1.454% | 2.857% | 8.934% | -46.7% |  | 7.064% |  | 0.3754% | 2.857% | -10.7% | -1.819% | -70.08% |  | 4.259%
| Z-Score (Absolute) |  | 0.72 | 1.84 | 0.84 | 0.35 | 0.07 | 0 | 0.3 | 2.46 | 1.86 | 0.21 | 1.86 | 0.12 | 0 | 0.67 | NA | 3.63 | 1.86 | 0.07
Results from laboratory 2.16 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 290 (Certified Value 99 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 100.2232 | 97 | 94 | 106 | 93.8 | 95.8 | 97.5 | 106.2 | 108 | 97 | 100 | 99.164 | 100.99 | 100 | 99.8 | 95.2 | 83.308 | 98.7 | 1
| 1 | 2 | 93.7885 | 95 | 93 | 105.5 | 97.5 | 100.5 | 97.2 | 100 | 107 | 97.1 | 99.9 | 99.4799 | 98.18 | 104 | 103 | 93.7 | 82.467 | 104 | 7
| 1 | 3 | 99.3444 | 103 | 95 | 108 | 91.3 | 100.5 | 97.4 | 97.9 | 106 | 100.4 | 102 | 99.7935 | 96.94 | 102 | 101.9 | 91.4 | 84.1132 | 93.6 | 16
| 1 | 4 | 93.7414 | 98 | 96 | 104.5 | 98.7 | 103.5 | 98.5 | 91.1 | 106 | 95.1 | 100 | 100.2592 | 98.45 | 101 | 99.7 | 96 | 83.062 | 91.1 | 99.1289
| 1 | 5 | 94.1546 | 92 | 98 | 104 | 91.1 | 95.3 | 98.2 | 81.2 | 107 | 100.1 | 101 | 100.8291 | 99.44 | 102 | 102.8 | 93.2 | 82.6965 | 105 | 58
| 1 | 6 | 102.8167 | 98 | 97 | 105.5 | 92.1 | 101.5 | 96.1 | 84.9 | 109 | 101.9 | 101 | 101.1444 | 97.99 | 104 | 102.6 | 93.1 | 82.584 | 106
| Mean |  | 97.3448 | 97.1667 | 95.5 | 105.5833 | 94.0833 | 99.5167 | 97.4833 | 93.55 | 107.1667 | 98.6 | 100.65 | 100.1117 | 98.665 | 102.1667 | 101.6333 | 93.7667 | 83.0385 | 99.7333
| Median |  | 96.7495 | 97.5 | 95.5 | 105.5 | 92.95 | 100.5 | 97.45 | 94.5 | 107 | 98.6 | 100.5 | 100.0263 | 98.315 | 102 | 102.25 | 93.45 | 82.8792 | 101.35
| Std Dev. |  | 3.9506 | 3.656 | 1.8708 | 1.3934 | 3.2756 | 3.2658 | 0.84242 | 9.5278 | 1.169 | 2.5861 | 0.83367 | 0.77479 | 1.3942 | 1.6021 | 1.5055 | 1.6379 | 0.6129 | 6.2997
| Rel.Std.Dev. |  | 4.058% | 3.763% | 1.959% | 1.32% | 3.482% | 3.282% | 0.8642% | 10.18% | 1.091% | 2.623% | 0.8283% | 0.7739% | 1.413% | 1.568% | 1.481% | 1.747% | 0.7381% | 6.317%
| PDM3 |  | -1.8% | -1.979% | -3.661% | 6.511% | -5.09% | 0.3912% | -1.66% | -5.628% | 8.108% | -0.5336% | 1.534% | 0.9914% | -0.468% | 3.064% | 2.526% | -5.409% | -16.23% | 0.6097%
| Z-Score (Absolute) |  | 0.35 | 0.39 | 0.84 | 1.87 | 1.22 | 0.24 | 0.31 | 1.37 | 2.29 | 0.01 | 0.54 | 0.4 | 0.01 | 0.95 | 0.81 | 1.31 | 4.19 | 0.3
| Analytical results for Ta in OREAS 290 (Certified Value 0.96 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.4 | 0.5 | 1.03 | 1 | 1.03 | 1.06 | < 0.05 | 1.11 | 1.0285 | 1.04 | 0.64 | 0.6 | < 0.05 | < 5 | 1.24 |  |  |  | 1
| 1 | 2 | 0.3 | 0.5 | 0.96 | 1.01 | 1.05 | 1.06 | 0.11 | 1.1 | 1.0321 | 1 | 0.64 | 0.7 | < 0.05 | < 5 | 1.32 |  |  |  | 35
| 1 | 3 | 0.4 | 0.7 | 0.99 | 0.95 | 1.06 | 1.07 | 0.1 | 1.1 | 1.0449 | 1.01 | 0.56 | 0.6 | < 0.05 | < 5 | 1.01 |  |  |  | 16
| 1 | 4 | 0.4 | 0.7 | 0.98 | 0.93 | 1.07 | 1.06 | 0.18 | 1.09 | 1.0477 | 1.05 | 0.46 | 0.6 | < 0.05 | < 5 | 1.17 |  |  |  | 0.9575
| 1 | 5 | 0.2 | 0.6 | 1 | 0.9 | 0.96 | 1.02 | 0.11 | 1.1 | 1.0476 | 1.01 | 0.46 | 0.5 | < 0.05 | < 5 | 1.24 |  |  |  | 59
| 1 | 6 | 0.3 | 0.6 | 0.98 | 0.95 | 1.08 | 1.04 | 0.14 | 1 | 1.031 | 0.98 | 0.47 | 0.6 | < 0.05 | < 5 | 1.01
| Mean |  | 0.33333 | 0.6 | 0.99 | 0.95667 | 1.0417 | 1.0517 | 0.128 | 1.0833 | 1.0386 | 1.015 | 0.53833 | 0.6 |  |  | 1.165
| Median |  | 0.35 | 0.6 | 0.985 | 0.95 | 1.055 | 1.06 | 0.11 | 1.1 | 1.0385 | 1.01 | 0.515 | 0.6 |  |  | 1.205
| Std Dev. |  | 0.08165 | 0.089443 | 0.023664 | 0.041793 | 0.043551 | 0.018348 | 0.032711 | 0.041312 | 0.0090028 | 0.025884 | 0.087274 | 0.063246 |  |  | 0.12911
| Rel.Std.Dev. |  | 24.49% | 14.91% | 2.39% | 4.369% | 4.181% | 1.745% | 25.56% | 3.813% | 0.8668% | 2.55% | 16.21% | 10.54% |  |  | 11.08%
| PDM3 |  | -65.19% | -37.34% | 3.395% | -0.08677% | 8.791% | 9.835% | -86.63% | 13.14% | 8.475% | 6.005% | -43.78% | -37.34% |  |  | 21.67%
| Z-Score (Absolute) |  | 2.53 | 1.5 | 0 | 0.13 | 0.2 | 0.24 | 3.39 | 0.36 | 0.19 | 0.1 | 1.74 | 1.5 | 3.72 | 5.82 | 0.67
| Analytical results for Tb in OREAS 290 (Certified Value 0.55 ppm) | [19 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.03 | 2.04 | 2.09 | 2.11 | 2.14 | 2.16 | 2.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.5 | 0.5 | 0.64 | 0.4 | 0.65 | 0.54 | 0.66 | [11 empty] | 1
| 1 | 2 | 0.6 | 0.5 | 0.49 | 0.48 | 0.65 | 0.52 | 0.65 | [11 empty] | 19
| 1 | 3 | 0.7 | 0.5 | 0.51 | 0.47 | 0.63 | 0.51 | 0.59 | [11 empty] | 16
| 1 | 4 | 0.7 | 0.5 | 0.41 | 0.45 | 0.64 | 0.53 | 0.62 | [11 empty] | 0.55343
| 1 | 5 | 0.5 | 0.5 | 0.43 | 0.47 | 0.63 | 0.5 | 0.66 | [11 empty] | 60
| 1 | 6 | 0.7 | 0.5 | 0.42 | 0.5 | 0.65 | 0.53 | 0.64
| Mean |  | 0.61667 | 0.5 | 0.48333 | 0.46167 | 0.64167 | 0.52167 | 0.63667
| Median |  | 0.65 | 0.5 | 0.46 | 0.47 | 0.645 | 0.525 | 0.645
| Std Dev. |  | 0.098319 | 0 | 0.086641 | 0.034303 | 0.0098319 | 0.01472 | 0.027325
| Rel.Std.Dev. |  | 15.94% | 0% | 17.93% | 7.43% | 1.532% | 2.822% | 4.292%
| PDM3 |  | 11.43% | -9.654% | -12.67% | -16.58% | 15.94% | -5.739% | 15.04%
| Z-Score (Absolute) |  | 1.07 | 0.24 | 0.43 | 0.67 | 1.35 | 0 | 1.29
| Analytical results for Te in OREAS 290 (Certified Value 0.072 ppm) | [19 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | < 0.05 | 0.05 | 0.07 | 0.08 | 0.07 | 0.09 | < 0.05 | 0.05 | < 0.1 | < 0.2 | < 0.1 | < 0.5 | < 0.05 | < 5 | 0.07 |  |  |  | 1
| 1 | 2 | < 0.05 | 0.06 | 0.07 | 0.11 | 0.07 | 0.08 | < 0.05 | 0.09 | < 0.1 | < 0.2 | 0.1 | < 0.5 | < 0.05 | < 5 | 0.05 |  |  |  | 37
| 1 | 3 | < 0.05 | 0.07 | 0.09 | 0.12 | 0.05 | 0.08 | < 0.05 | 0.06 | < 0.1 | < 0.2 | < 0.1 | < 0.5 | < 0.05 | < 5 | 0.06 |  |  |  | 16
| 1 | 4 | < 0.05 | 0.07 | 0.06 | 0.06 | 0.07 | 0.07 | < 0.05 | 0.05 | < 0.1 | < 0.2 | < 0.1 | < 0.5 | < 0.05 | < 5 | 0.05 |  |  |  | 0.071667
| 1 | 5 | < 0.05 | 0.06 | 0.07 | 0.13 | 0.05 | 0.07 | < 0.05 | 0.06 | < 0.1 | < 0.2 | < 0.1 | < 0.5 | < 0.05 | < 5 | 0.07 |  |  |  | 61
| 1 | 6 | < 0.05 | 0.06 | 0.09 | 0.06 | 0.06 | 0.07 | < 0.05 | 0.05 | < 0.1 | < 0.2 | < 0.1 | < 0.5 | < 0.05 | < 5 | 0.05
| Mean |  |  | 0.061667 | 0.075 | 0.093333 | 0.061667 | 0.076667 |  | 0.06 |  |  | 0.1 |  |  |  | 0.058333
| Median |  |  | 0.06 | 0.07 | 0.095 | 0.065 | 0.075 |  | 0.055 |  |  | 0.1 |  |  |  | 0.055
| Std Dev. |  |  | 0.0075277 | 0.012247 | 0.030768 | 0.0098319 | 0.008165 |  | 0.015492 | [6 empty] | 0.0098319
| Rel.Std.Dev. |  |  | 12.21% | 16.33% | 32.97% | 15.94% | 10.65% |  | 25.82% | [6 empty] | 16.85%
| PDM3 |  |  | -13.95% | 4.651% | 30.23% | -13.95% | 6.977% |  | -16.28% |  |  | 39.53% |  |  |  | -18.6%
| Z-Score (Absolute) |  | 1.65 | 0 | 0.6 | 1.42 | 0 | 0.67 | 1.65 | 0.07 | 0.52 | 1.72 | 0.15 | 8.47 | 1.65 | 109.61 | 0.15
| Analytical results for Th in OREAS 290 (Certified Value 15.1 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 11 | 14.2 | 17.1 | 14.15 | 16.45 | 14.45 | 13.2 | 10 | 15.7101 | 14.88 | 15.4 | 13 | 14.6 | 16.4 |  |  |  |  | 1
| 1 | 2 | 13.2 | 13.8 | 16.7 | 14.75 | 16.85 | 14.8 | 11.4 | 11 | 15.6603 | 14.88 | 16.6 | 13.7 | 15.4 | 16 |  |  |  |  | 38
| 1 | 3 | 15.2 | 14.4 | 17.4 | 13.55 | 17.15 | 15.65 | 12.4 | 10.6 | 15.5449 | 14.79 | 15.8 | 13.4 | 15 | 14.6 |  |  |  |  | 16
| 1 | 4 | 13.7 | 14.3 | 16.3 | 15.85 | 16.8 | 15.05 | 9.6 | 10.1 | 15.5852 | 15.11 | 15.6 | 13.5 | 16.3 | 15.7 |  |  |  |  | 15.1315
| 1 | 5 | 12.3 | 14.5 | 17.45 | 13.15 | 15.65 | 15.45 | 10.1 | 11.7 | 15.7505 | 14.76 | 16.3 | 14 | 14.9 | 15.9 |  |  |  |  | 62
| 1 | 6 | 12.9 | 13.6 | 16.6 | 13.65 | 16.85 | 14.7 | 9.9 | 11.3 | 15.5106 | 14.76 | 16.3 | 13.6 | 15.8 | 15.5
| Mean |  | 13.05 | 14.1333 | 16.925 | 14.1833 | 16.625 | 15.0167 | 11.1 | 10.7833 | 15.6269 | 14.8633 | 16 | 13.5333 | 15.3333 | 15.6833
| Median |  | 13.05 | 14.25 | 16.9 | 13.9 | 16.825 | 14.925 | 10.75 | 10.8 | 15.6227 | 14.835 | 16.05 | 13.55 | 15.2 | 15.8
| Std Dev. |  | 1.4039 | 0.3559 | 0.46449 | 0.98522 | 0.52702 | 0.46007 | 1.4751 | 0.67355 | 0.095194 | 0.13277 | 0.46904 | 0.33267 | 0.6314 | 0.61128
| Rel.Std.Dev. |  | 10.76% | 2.518% | 2.744% | 6.946% | 3.17% | 3.064% | 13.29% | 6.246% | 0.6092% | 0.8932% | 2.932% | 2.458% | 4.118% | 3.898%
| PDM3 |  | -13.76% | -6.597% | 11.85% | -6.266% | 9.87% | -0.7592% | -26.64% | -28.74% | 3.274% | -1.773% | 5.739% | -10.56% | 1.334% | 3.647%
| Z-Score (Absolute) |  | 1.37 | 0.58 | 1.43 | 0.55 | 1.22 | 0.06 | 2.77 | 3 | 0.5 | 0.06 | 0.77 | 1.02 | 0.28 | 0.54
| Analytical results for Ti in OREAS 290 (Certified Value 0.359 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.36259 | 0.3286 | 0.347 | 0.386 | 0.366 | 0.378 | 0.372 | 0.32 | 0.36 | 0.33 | 0.35 | 0.37465 | 0.3723 | 0.371 | 0.2876 | 0.27 | 0.399 | 0.388 | 1
| 1 | 2 | 0.33289 | 0.3401 | 0.34 | 0.392 | 0.368 | 0.393 | 0.375 | 0.31 | 0.3 | 0.33 | 0.36 | 0.38942 | 0.3792 | 0.376 | 0.2996 | 0.272 | 0.399 | 0.388 | 8
| 1 | 3 | 0.35698 | 0.3526 | 0.358 | 0.398 | 0.362 | 0.395 | 0.384 | 0.32 | 0.33 | 0.33 | 0.37 | 0.39058 | 0.3783 | 0.311 | 0.2958 | 0.258 | 0.399 | 0.368 | 16
| 1 | 4 | 0.33172 | 0.3505 | 0.365 | 0.392 | 0.381 | 0.393 | 0.381 | 0.32 | 0.34 | 0.33 | 0.38 | 0.38525 | 0.381 | 0.31 | 0.2926 | 0.236 | 0.399 | 0.33 | 0.35905
| 1 | 5 | 0.33433 | 0.3286 | 0.361 | 0.393 | 0.374 | 0.375 | 0.381 | 0.32 | 0.33 | 0.34 | 0.37 | 0.39388 | 0.3886 | 0.351 | 0.3005 | 0.211 | 0.3885 | 0.401 | 63
| 1 | 6 | 0.35806 | 0.3432 | 0.354 | 0.395 | 0.363 | 0.402 | 0.376 | 0.33 | 0.34 | 0.34 | 0.37 | 0.37766 | 0.3789 | 0.33 | 0.2983 | 0.251 | 0.378 | 0.413
| Mean |  | 0.3461 | 0.3406 | 0.35417 | 0.39267 | 0.369 | 0.38933 | 0.37817 | 0.32 | 0.33333 | 0.33333 | 0.36667 | 0.38524 | 0.37972 | 0.3415 | 0.29573 | 0.24967 | 0.39375 | 0.38133
| Median |  | 0.34566 | 0.34165 | 0.356 | 0.3925 | 0.367 | 0.393 | 0.3785 | 0.32 | 0.335 | 0.33 | 0.37 | 0.38734 | 0.37905 | 0.3405 | 0.29705 | 0.2545 | 0.399 | 0.388
| Std Dev. |  | 0.014514 | 0.010364 | 0.0092826 | 0.0039833 | 0.0072664 | 0.01052 | 0.004535 | 0.0063246 | 0.019664 | 0.005164 | 0.010328 | 0.0076202 | 0.0052632 | 0.029002 | 0.0049037 | 0.023088 | 0.0087849 | 0.029296
| Rel.Std.Dev. |  | 4.194% | 3.043% | 2.621% | 1.014% | 1.969% | 2.702% | 1.199% | 1.976% | 5.899% | 1.549% | 2.817% | 1.978% | 1.386% | 8.492% | 1.658% | 9.248% | 2.231% | 7.683%
| PDM3 |  | -3.607% | -5.138% | -1.359% | 9.364% | 2.772% | 8.435% | 5.325% | -10.87% | -7.161% | -7.161% | 2.122% | 7.296% | 5.757% | -4.887% | -17.63% | -30.46% | 9.666% | 6.207%
| Z-Score (Absolute) |  | 0.42 | 0.58 | 0.18 | 0.95 | 0.25 | 0.85 | 0.52 | 1.19 | 0.8 | 0.8 | 0.18 | 0.73 | 0.57 | 0.56 | 1.91 | 3.27 | 0.98 | 0.62
| Analytical results for Tl in OREAS 290 (Certified Value 0.86 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.86 | 0.82 | 0.9 | 0.77 | 0.8 | 0.84 | 0.83 | 0.8 | 0.87225 | 0.87 | 0.81 | 0.9 | 0.9 | 0.97 |  |  |  |  | 1
| 1 | 2 | 0.83 | 0.81 | 0.82 | 0.79 | 0.82 | 0.83 | 0.87 | 0.88 | 0.8765 | 0.86 | 0.85 | 1 | 0.93 | 0.95 |  |  |  |  | 40
| 1 | 3 | 0.81 | 0.83 | 0.87 | 0.79 | 0.8 | 0.89 | 0.88 | 0.83 | 0.88405 | 0.9 | 0.85 | 0.9 | 0.91 | 0.81 |  |  |  |  | 16
| 1 | 4 | 0.82 | 0.81 | 0.85 | 0.8 | 0.83 | 0.86 | 0.87 | 0.84 | 0.85817 | 0.92 | 0.84 | 0.9 | 0.94 | 0.92 |  |  |  |  | 0.85951
| 1 | 5 | 0.82 | 0.84 | 0.88 | 0.71 | 0.75 | 0.87 | 0.87 | 0.86 | 0.88535 | 0.87 | 0.9 | 0.9 | 0.93 | 0.92 |  |  |  |  | 64
| 1 | 6 | 0.91 | 0.79 | 0.86 | 0.75 | 0.81 | 0.86 | 0.89 | 0.79 | 0.86487 | 0.86 | 0.84 | 1 | 0.9 | 0.91
| Mean |  | 0.84167 | 0.81667 | 0.86333 | 0.76833 | 0.80167 | 0.85833 | 0.86833 | 0.83333 | 0.87353 | 0.88 | 0.84833 | 0.93333 | 0.91833 | 0.91333
| Median |  | 0.825 | 0.815 | 0.865 | 0.78 | 0.805 | 0.86 | 0.87 | 0.835 | 0.87438 | 0.87 | 0.845 | 0.9 | 0.92 | 0.92
| Std Dev. |  | 0.037639 | 0.017512 | 0.027325 | 0.033714 | 0.027869 | 0.02137 | 0.020412 | 0.034448 | 0.010694 | 0.024495 | 0.029269 | 0.05164 | 0.017224 | 0.055377
| Rel.Std.Dev. |  | 4.472% | 2.144% | 3.165% | 4.388% | 3.476% | 2.49% | 2.351% | 4.134% | 1.224% | 2.784% | 3.45% | 5.533% | 1.876% | 6.063%
| PDM3 |  | -2.076% | -4.985% | 0.4443% | -10.61% | -6.73% | -0.1374% | 1.026% | -3.046% | 1.631% | 2.383% | -1.301% | 8.588% | 6.843% | 6.262%
| Z-Score (Absolute) |  | 0.55 | 1.28 | 0.07 | 2.67 | 1.71 | 0.07 | 0.22 | 0.79 | 0.37 | 0.55 | 0.36 | 2.1 | 1.66 | 1.52
| Analytical results for Tm in OREAS 290 (Certified Value 0.22 ppm) | [19 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.14 | 2.16 | 2.20 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [13 empty] | 2
2
| 1 | 1 | 0.2 | 0.2 | 0.28 | 0.2 | 0.24 | [13 empty] | 1
| 1 | 2 | 0.2 | 0.2 | 0.27 | 0.2 | 0.25 | [13 empty] | 20
| 1 | 3 | 0.2 | 0.2 | 0.27 | 0.2 | 0.23 | [13 empty] | 16
| 1 | 4 | 0.2 | 0.2 | 0.28 | 0.2 | 0.23 | [13 empty] | 0.223
| 1 | 5 | 0.2 | 0.2 | 0.27 | 0.2 | 0.26 | [13 empty] | 65
| 1 | 6 | 0.2 | 0.2 | 0.26 | 0.2 | 0.25
| Mean |  | 0.2 | 0.2 | 0.27167 | 0.2 | 0.24333
| Median |  | 0.2 | 0.2 | 0.27 | 0.2 | 0.245
| Std Dev. |  | 3.0405e-17 | 3.0405e-17 | 0.0075277 | 3.0405e-17 | 0.012111
| Rel.Std.Dev. |  | 1.52e-14% | 1.52e-14% | 2.771% | 1.52e-14% | 4.977%
| PDM3 |  | -10.31% | -10.31% | 21.82% | -10.31% | 9.118%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for U in OREAS 290 (Certified Value 2.79 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 3 | 2.6 | 2.8 | 2.6 | 2.5 | 2.5 | 3.1 | 2.4 | 2.9897 | 2.86 | 2.74 | 2.7 | 2.8 | 3.04 |  |  |  |  | 1
| 1 | 2 | 3.1 | 2.5 | 2.9 | 2.7 | 2.6 | 2.6 | 2.8 | 2.7 | 3.164 | 2.86 | 2.85 | 2.8 | 2.9 | 3 |  |  |  |  | 42
| 1 | 3 | 2.9 | 2.6 | 2.9 | 2.5 | 2.5 | 2.7 | 2.9 | 2.6 | 2.953 | 2.88 | 2.64 | 2.8 | 2.9 | 2.66 |  |  |  |  | 16
| 1 | 4 | 3.2 | 2.5 | 2.8 | 2.8 | 2.6 | 2.6 | 2.7 | 2.5 | 3.1572 | 2.91 | 2.66 | 2.8 | 3 | 2.93 |  |  |  |  | 2.785
| 1 | 5 | 3.2 | 2.6 | 2.8 | 2.5 | 2.4 | 2.7 | 2.7 | 2.8 | 2.9986 | 2.9 | 2.8 | 2.9 | 2.9 | 2.99 |  |  |  |  | 66
| 1 | 6 | 3.4 | 2.5 | 2.8 | 2.6 | 2.6 | 2.6 | 2.8 | 2.7 | 2.9327 | 2.85 | 2.75 | 2.7 | 2.9 | 2.93
| Mean |  | 3.1333 | 2.55 | 2.8333 | 2.6167 | 2.5333 | 2.6167 | 2.8333 | 2.6167 | 3.0325 | 2.8767 | 2.74 | 2.7833 | 2.9 | 2.925
| Median |  | 3.15 | 2.55 | 2.8 | 2.6 | 2.55 | 2.6 | 2.8 | 2.65 | 2.9942 | 2.87 | 2.745 | 2.8 | 2.9 | 2.96
| Std Dev. |  | 0.17512 | 0.054772 | 0.05164 | 0.1169 | 0.08165 | 0.075277 | 0.15055 | 0.1472 | 0.10207 | 0.024221 | 0.08025 | 0.075277 | 0.063246 | 0.13664
| Rel.Std.Dev. |  | 5.589% | 2.148% | 1.823% | 4.468% | 3.223% | 2.877% | 5.314% | 5.625% | 3.366% | 0.842% | 2.929% | 2.705% | 2.181% | 4.671%
| PDM3 |  | 12.51% | -8.439% | 1.734% | -6.046% | -9.038% | -6.046% | 1.734% | -6.046% | 8.886% | 3.29% | -1.617% | -0.06119% | 4.128% | 5.026%
| Z-Score (Absolute) |  | 1.42 | 1.13 | 0.11 | 0.84 | 1.2 | 0.84 | 0.11 | 0.84 | 0.98 | 0.3 | 0.3 | 0.11 | 0.4 | 0.51
| Analytical results for V in OREAS 290 (Certified Value 92 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 99.8342 | 82 | 91 | 97 | 89 | 88 | 92 | 64 | 97 | 90 | 87 | 87.4511 | 92 | 85 | 103 | 65 | 92.1202 | 95 | 1
| 1 | 2 | 88.6547 | 84 | 91 | 98 | 91 | 93 | 95 | 76 | 85 | 88 | 89 | 87.5667 | 91 | 95 | 104 | 63 | 92.6051 | 100 | 9
| 1 | 3 | 97.6424 | 87 | 94 | 101 | 88 | 95 | 97 | 78 | 93 | 89 | 86 | 89.2082 | 93 | 81 | 101 | 59 | 92.3008 | 91 | 16
| 1 | 4 | 89.7415 | 87 | 96 | 98 | 96 | 92 | 96 | 81 | 91 | 91 | 86 | 86.7056 | 95 | 80 | 101 | 53 | 92.4572 | 86 | 91.982
| 1 | 5 | 88.4997 | 80 | 94 | 98 | 92 | 89 | 92 | 79 | 89 | 92 | 88 | 89.8177 | 90 | 90 | 103 | 57 | 91.7965 | 98 | 67
| 1 | 6 | 94.9643 | 85 | 95 | 100 | 89 | 95 | 95 | 82 | 93 | 94 | 86 | 85.8088 | 92 | 86 | 100 | 52 | 93.0982 | 102
| Mean |  | 93.2228 | 84.1667 | 93.5 | 98.6667 | 90.8333 | 92 | 94.5 | 76.6667 | 91.3333 | 90.6667 | 87 | 87.7597 | 92.1667 | 86.1667 | 102 | 58.1667 | 92.3963 | 95.3333
| Median |  | 92.3529 | 84.5 | 94 | 98 | 90 | 92.5 | 95 | 78.5 | 92 | 90.5 | 86.5 | 87.5089 | 92 | 85.5 | 102 | 58 | 92.379 | 96.5
| Std Dev. |  | 4.9309 | 2.7869 | 2.0736 | 1.5055 | 2.9269 | 2.9665 | 2.0736 | 6.5625 | 4.0825 | 2.1602 | 1.2649 | 1.5089 | 1.7224 | 5.6362 | 1.5492 | 5.2313 | 0.44389 | 5.9889
| Rel.Std.Dev. |  | 5.289% | 3.311% | 2.218% | 1.526% | 3.222% | 3.224% | 2.194% | 8.56% | 4.47% | 2.383% | 1.454% | 1.719% | 1.869% | 6.541% | 1.519% | 8.994% | 0.4804% | 6.282%
| PDM3 |  | 1.349% | -8.497% | 1.65% | 7.267% | -1.249% | 0.01956% | 2.737% | -16.65% | -0.7052% | -1.43% | -5.416% | -4.59% | 0.2008% | -6.322% | 10.89% | -36.76% | 0.4505% | 3.643%
| Z-Score (Absolute) |  | 0.32 | 1.56 | 0.38 | 1.45 | 0.17 | 0.07 | 0.59 | 3.11 | 0.07 | 0.21 | 0.97 | 0.81 | 0.1 | 1.14 | 2.14 | 6.95 | 0.15 | 0.76
| Analytical results for W in OREAS 290 (Certified Value 2.83 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 2
2
| 1 | 1 | < 4 | 0.4 | 0.6 | 3.3 | 3.2 | 3 | 3.1 | < 0.1 | < 10 | 2 | < 10 | 2.8363 | 2.7 | 3.3 | < 50 | < 0.1 | < 10 | 2.7 | 1
| 1 | 2 | < 4 | 0.4 | 0.6 | 3 | 3.1 | 3.1 | 3 | 0.2 | < 10 | 2.2 | < 10 | 2.8145 | 2.9 | 3 | < 50 | < 0.1 | < 10 | 2.6 | 10
| 1 | 3 | < 4 | 0.4 | 1.2 | 3 | 3.1 | 3.1 | 3.2 | 0.2 | < 10 | 2.2 | < 10 | 2.7612 | 3 | 0.6 | < 50 | < 0.1 | < 10 | 2.3 | 16
| 1 | 4 | < 4 | 0.4 | 1.4 | 3 | 2.8 | 3 | 3.2 | 0.2 | < 10 | 2.2 | < 10 | 2.9235 | 3.1 | 0.5 | < 50 | 0.1 | < 10 | 2.2 | 2.8289
| 1 | 5 | < 4 | 0.3 | 0.8 | 3.1 | 2.8 | 2.9 | 3.1 | 0.2 | < 10 | 2.3 | < 10 | 2.9044 | 3 | 0.5 | < 50 | < 0.1 | < 10 | 2.3 | 68
| 1 | 6 | < 4 | 0.5 | 0.8 | 3.2 | 3 | 3.1 | 3.1 | 0.3 | < 10 | 2.2 | < 10 | 2.7452 | 2.8 | 0.5 | < 50 | < 0.1 | < 10 | 2.6
| Mean |  |  | 0.4 | 0.9 | 3.1 | 3 | 3.0333 | 3.1167 | 0.22 |  | 2.1833 |  | 2.8308 | 2.9167 | 1.4 |  | 0.1 |  | 2.45
| Median |  |  | 0.4 | 0.8 | 3.05 | 3.05 | 3.05 | 3.1 | 0.2 |  | 2.2 |  | 2.8254 | 2.95 | 0.55 |  | 0.1 |  | 2.45
| Std Dev. |  |  | 0.063246 | 0.32863 | 0.12649 | 0.16733 | 0.08165 | 0.075277 | 0.044721 |  | 0.098319 |  | 0.072776 | 0.1472 | 1.3594 |  |  |  | 0.20736
| Rel.Std.Dev. |  |  | 15.81% | 36.51% | 4.08% | 5.578% | 2.692% | 2.415% | 20.33% |  | 4.503% |  | 2.571% | 5.047% | 97.1% |  |  |  | 8.464%
| PDM3 |  |  | -85.86% | -68.18% | 9.585% | 6.05% | 7.228% | 10.17% | -92.22% |  | -22.82% |  | 0.06995% | 3.104% | -50.51% |  | -96.46% |  | -13.39%
| Z-Score (Absolute) |  | 0.75 | 2.13 | 1.7 | 0.2 | 0.11 | 0.14 | 0.21 | 2.31 | 1.83 | 0.6 | 1.83 | 0.04 | 0.04 | NA | NA | 2.43 | 1.83 | 0.37
| Analytical results for Y in OREAS 290 (Certified Value 15.2 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 1
1
| 1 | 1 | 13.8 | 13.1 | 13.8 | 17 | 14.6 | 16.4 | 14.8 | 15.1 | 15.6 | 16.2 | 13.4 | 16.8784 | 15.68 | 17.6 | 12.2 | 17.4 | 29.263 | 15.4 | 1
| 1 | 2 | 13.225 | 13.6 | 13.6 | 17.1 | 17.8 | 17.1 | 15.1 | 12.3 | 14.9 | 16.3 | 13.5 | 17.4158 | 15.35 | 18.9 | 12.4 | 18 | 29.3861 | 15.6 | 11
| 1 | 3 | 13.685 | 15.1 | 13.3 | 17.6 | 14.3 | 16.8 | 15.2 | 13.3 | 14 | 16.3 | 13.5 | 16.7596 | 15.08 | 18.2 | 12.6 | 17.6 | 29.65 | 14.6 | 16
| 1 | 4 | 13.225 | 14.3 | 13.5 | 17.1 | 15.7 | 17.6 | 15.4 | 10.2 | 15 | 16.7 | 13.4 | 17.0875 | 15.64 | 16.6 | 12.5 | 20.6 | 29.3131 | 14.9 | 15.1844
| 1 | 5 | 12.65 | 13.2 | 14.1 | 16.5 | 12.7 | 15.8 | 15.2 | 10.5 | 13.1 | 16.9 | 13.4 | 17.5835 | 15.76 | 17.2 | 12.3 | 19.1 | 29.4567 | 15.9 | 69
| 1 | 6 | 14.145 | 13.5 | 13.6 | 16.8 | 13.9 | 17 | 14.8 | 10.6 | 14.9 | 16.8 | 13.7 | 16.6295 | 15.61 | 17.5 | 12.1 | 19 | 29.541 | 15.5
| Mean |  | 13.455 | 13.8 | 13.65 | 17.0167 | 14.8333 | 16.7833 | 15.0833 | 12 | 14.5833 | 16.5333 | 13.4833 | 17.0591 | 15.52 | 17.6667 | 12.35 | 18.6167 | 29.435 | 15.3167
| Median |  | 13.455 | 13.55 | 13.6 | 17.05 | 14.45 | 16.9 | 15.15 | 11.45 | 14.9 | 16.5 | 13.45 | 16.9829 | 15.625 | 17.55 | 12.35 | 18.5 | 29.4214 | 15.45
| Std Dev. |  | 0.5295 | 0.7642 | 0.27386 | 0.3656 | 1.75 | 0.62102 | 0.24014 | 1.9411 | 0.88863 | 0.30111 | 0.1169 | 0.37682 | 0.2562 | 0.79917 | 0.18708 | 1.2007 | 0.14477 | 0.47924
| Rel.Std.Dev. |  | 3.935% | 5.538% | 2.006% | 2.149% | 11.8% | 3.7% | 1.592% | 16.18% | 6.093% | 1.821% | 0.867% | 2.209% | 1.651% | 4.524% | 1.515% | 6.45% | 0.4918% | 3.129%
| PDM3 |  | -11.39% | -9.117% | -10.11% | 12.07% | -2.312% | 10.53% | -0.6658% | -20.97% | -3.959% | 8.883% | -11.2% | 12.35% | 2.21% | 16.35% | -18.67% | 22.6% | 93.85% | 0.8708%
| Z-Score (Absolute) |  | 0.71 | 0.57 | 0.63 | 0.74 | 0.15 | 0.65 | 0.05 | 1.31 | 0.25 | 0.54 | 0.7 | 0.76 | 0.13 | 1.01 | 1.16 | 1.4 | 5.82 | 0.05
| Analytical results for Yb in OREAS 290 (Certified Value 1.58 ppm) | [19 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.03 | 2.04 | 2.09 | 2.11 | 2.14 | 2.16 | 2.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 1.5 | 1.5 | 1.7 | 1.5 | 1.89 | 1.5 | 1.65 | [11 empty] | 1
| 1 | 2 | 1.7 | 1.4 | 1.5 | 1.6 | 1.87 | 1.4 | 1.69 | [11 empty] | 21
| 1 | 3 | 1.8 | 1.5 | 1.5 | 1.5 | 1.86 | 1.5 | 1.58 | [11 empty] | 16
| 1 | 4 | 1.6 | 1.4 | 1.2 | 1.5 | 1.88 | 1.4 | 1.65 | [11 empty] | 1.5844
| 1 | 5 | 1.6 | 1.5 | 1.3 | 1.7 | 1.88 | 1.6 | 1.65 | [11 empty] | 70
| 1 | 6 | 1.7 | 1.4 | 1.2 | 1.7 | 1.83 | 1.5 | 1.64
| Mean |  | 1.65 | 1.45 | 1.4 | 1.5833 | 1.8683 | 1.4833 | 1.6433
| Median |  | 1.65 | 1.45 | 1.4 | 1.55 | 1.875 | 1.5 | 1.65
| Std Dev. |  | 0.10488 | 0.054772 | 0.2 | 0.098319 | 0.02137 | 0.075277 | 0.03559
| Rel.Std.Dev. |  | 6.356% | 3.777% | 14.29% | 6.21% | 1.144% | 5.075% | 2.166%
| PDM3 |  | 4.138% | -8.484% | -11.64% | -0.06913% | 17.92% | -6.381% | 3.718%
| Z-Score (Absolute) |  | 0.45 | 0.9 | 1.24 | 0 | 1.92 | 0.67 | 0.4
| Analytical results for Zn in OREAS 290 (Certified Value 98 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 110.21 | 94.7 | 92 | 102 | 92 | 92 | 96 | 95 | 91 | 88 | 92 | 101.6451 | 102 | 92 | 98 | 92 | 102.391 | 100 | 1
| 1 | 2 | 103 | 95.9 | 90.7 | 103 | 93 | 96 | 98 | 75 | 92 | 90 | 92 | 100.7271 | 101 | 97 | 98 | 98 | 101.9106 | 95.5 | 12
| 1 | 3 | 110.21 | 93.9 | 92.4 | 106 | 90 | 97 | 99 | 76 | 90 | 86 | 93 | 104.1804 | 101 | 97 | 101 | 100 | 100.7275 | 97.1 | 16
| 1 | 4 | 105.06 | 95.8 | 93.5 | 103 | 96 | 94 | 98 | 71 | 93 | 89 | 95 | 101.6662 | 100 | 93 | 108 | 102 | 102.1533 | 102 | 97.5423
| 1 | 5 | 105.06 | 93.2 | 95.2 | 102 | 94 | 92 | 99 | 73 | 90 | 87 | 92 | 103.8054 | 99 | 102 | 101 | 101 | 100.812 | 97.4 | 71
| 1 | 6 | 112.27 | 94.8 | 92.5 | 103 | 91 | 100 | 97 | 75 | 89 | 91 | 94 | 102.7155 | 100 | 96 | 107 | 100 | 100.968 | 102
| Mean |  | 107.635 | 94.7167 | 92.7167 | 103.1667 | 92.6667 | 95.1667 | 97.8333 | 77.5 | 90.8333 | 88.5 | 93 | 102.4566 | 100.5 | 96.1667 | 102.1667 | 98.8333 | 101.4937 | 99
| Median |  | 107.635 | 94.75 | 92.45 | 103 | 92.5 | 95 | 98 | 75 | 90.5 | 88.5 | 92.5 | 102.1908 | 100.5 | 96.5 | 101 | 100 | 101.4393 | 98.7
| Std Dev. |  | 3.728 | 1.0534 | 1.5171 | 1.472 | 2.1602 | 3.1252 | 1.169 | 8.7579 | 1.472 | 1.8708 | 1.2649 | 1.3515 | 1.0488 | 3.5449 | 4.3551 | 3.6009 | 0.74057 | 2.7357
| Rel.Std.Dev. |  | 3.464% | 1.112% | 1.636% | 1.427% | 2.331% | 3.284% | 1.195% | 11.3% | 1.621% | 2.114% | 1.36% | 1.319% | 1.044% | 3.686% | 4.263% | 3.643% | 0.7297% | 2.763%
| PDM3 |  | 10.35% | -2.897% | -4.947% | 5.766% | -4.998% | -2.435% | 0.2984% | -20.55% | -6.878% | -9.27% | -4.657% | 5.038% | 3.032% | -1.41% | 4.741% | 1.324% | 4.051% | 1.494%
| Z-Score (Absolute) |  | 1.66 | 0.36 | 0.67 | 0.97 | 0.68 | 0.29 | 0.13 | 3.05 | 0.97 | 1.33 | 0.63 | 0.85 | 0.55 | 0.13 | 0.81 | 0.29 | 0.7 | 0.31
| Analytical results for Zr in OREAS 290 (Certified Value 133 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 124.3 | 134 | 119 | 144.5 | 130 | 141 | 125.5 | 131.5 | 138 | 140 | 115 | 142.1885 | 136 | 138 | 121 | 156 | 136.28 | 144 | 1
| 1 | 2 | 117.7 | 136.2 | 120.4 | 143.5 | 136 | 148 | 129.5 | 133.5 | 132 | 148 | 119 | 142.7229 | 133.7 | 141 | 126 | 145 | 135.59 | 143 | 48
| 1 | 3 | 124.3 | 135 | 121.9 | 148.5 | 125 | 144.5 | 129 | 132.7 | 123 | 141 | 129 | 145.122 | 134 | 130 | 122 | 138 | 135.06 | 133 | 16
| 1 | 4 | 117.7 | 141.5 | 124.4 | 143.5 | 136.5 | 149.5 | 131 | 133.6 | 133 | 146 | 123 | 142.4494 | 132.7 | 127 | 121 | 131 | 133.7694 | 136 | 133.3184
| 1 | 5 | 117.7 | 132.4 | 127.1 | 143 | 120.5 | 136 | 129.5 | 129.7 | 128 | 143 | 123 | 142.5924 | 138.3 | 134 | 120 | 128 | 133.1438 | 144 | 72
| 1 | 6 | 126.5 | 135.1 | 122.9 | 143.5 | 126 | 147.5 | 126.5 | 132.3 | 128 | 149 | 128 | 144.962 | 131.1 | 133 | 123 | 129 | 132.1675 | 140
| Mean |  | 121.3667 | 135.7 | 122.6167 | 144.4167 | 129 | 144.4167 | 128.5 | 132.2167 | 130.3333 | 144.5 | 122.8333 | 143.3395 | 134.3 | 133.8333 | 122.1667 | 137.8333 | 134.3351 | 140
| Median |  | 121 | 135.05 | 122.4 | 143.5 | 128 | 146 | 129.25 | 132.5 | 130 | 144.5 | 123 | 142.6577 | 133.85 | 133.5 | 121.5 | 134.5 | 134.4147 | 141.5
| Std Dev. |  | 4.0962 | 3.1151 | 2.8937 | 2.0595 | 6.3797 | 5.1129 | 2.0736 | 1.4593 | 5.164 | 3.7283 | 5.3072 | 1.3315 | 2.5354 | 5.1153 | 2.137 | 10.9803 | 1.5698 | 4.6043
| Rel.Std.Dev. |  | 3.375% | 2.296% | 2.36% | 1.426% | 4.945% | 3.54% | 1.614% | 1.104% | 3.962% | 2.58% | 4.321% | 0.9289% | 1.888% | 3.822% | 1.749% | 7.966% | 1.169% | 3.289%
| PDM3 |  | -8.965% | 1.786% | -8.027% | 8.325% | -3.239% | 8.325% | -3.614% | -0.8264% | -2.239% | 8.387% | -7.865% | 7.517% | 0.7363% | 0.3862% | -8.365% | 3.387% | 0.7626% | 5.012%
| Z-Score (Absolute) |  | 1.49 | 0.19 | 1.34 | 1.21 | 0.59 | 1.21 | 0.65 | 0.22 | 0.44 | 1.22 | 1.32 | 1.09 | 0.03 | 0.03 | 1.4 | 0.44 | 0.03 | 0.7
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C634: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C652: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C670: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C688: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C706: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C833: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C870: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C906: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C942: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C978: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C996: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1014: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1032: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1050: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1068: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1086: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1104: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 290 (Certified Value 0.192 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.2094 | 0.28 | 0.17 | 0.17 | 0.2 | 0.19 | 0.2 | 0.18 | < 2 | 0.19 | < 2 | 0.18033 | 0.19 | 0.39 | 0.19 | 0.181 | < 0.5 | < 0.5 | 0.18 | 1
| 1 | 2 | 0.1977 | 0.25 | 0.23 | 0.17 | 0.19 | 0.19 | 0.17 | 0.28 | < 2 | 0.18 | < 2 | 0.23061 | 0.19 | 0.38 | 0.18 | 0.2 | < 0.5 | < 0.5 | 0.18 | 32
| 1 | 3 | 0.2099 | 0.23 | 0.216 | 0.19 | 0.19 | 0.2 | 0.19 | 0.19 | < 2 | 0.17 | < 2 | 0.20449 | 0.19 | 0.37 | 0.18 | 0.193 | < 0.5 | < 0.5 | 0.16 | 16
| 1 | 4 | 0.2087 | 0.22 | 0.213 | 0.17 | 0.2 | 0.19 | 0.16 | 0.19 | < 2 | 0.17 | < 2 | 0.2233 | 0.2 | 0.37 | 0.21 | 0.211 | < 0.5 | < 0.5 | 0.18 | 0.19233
| 1 | 5 | 0.203 | 0.25 | 0.229 | 0.19 | 0.2 | 0.2 | 0.16 | 0.2 | < 2 | 0.17 | < 2 | 0.18905 | 0.2 | 0.38 | 0.23 | 0.199 | < 0.5 | < 0.5 | 0.18 | 74
| 1 | 6 | 0.2038 | 0.18 | 0.204 | 0.2 | 0.2 | 0.2 | 0.18 | 0.18 | < 2 | 0.18 | < 2 | 0.18152 | 0.21 | 0.37 | 0.21 | 0.196 | < 0.5 | < 0.5 | 0.17
| Mean |  | 0.20542 | 0.235 | 0.21033 | 0.18167 | 0.19667 | 0.195 | 0.17667 | 0.20333 |  | 0.17667 |  | 0.20155 | 0.19667 | 0.37667 | 0.2 | 0.19667 |  |  | 0.175
| Median |  | 0.20625 | 0.24 | 0.2145 | 0.18 | 0.2 | 0.195 | 0.175 | 0.19 |  | 0.175 |  | 0.19677 | 0.195 | 0.375 | 0.2 | 0.1975 |  |  | 0.18
| Std Dev. |  | 0.0047906 | 0.033912 | 0.022097 | 0.013292 | 0.005164 | 0.0054772 | 0.01633 | 0.038297 |  | 0.008165 |  | 0.021604 | 0.008165 | 0.008165 | 0.02 | 0.0098116 |  |  | 0.0083666
| Rel.Std.Dev. |  | 2.332% | 14.43% | 10.51% | 7.316% | 2.626% | 2.809% | 9.243% | 18.83% |  | 4.622% |  | 10.72% | 4.152% | 2.168% | 10% | 4.989% |  |  | 4.781%
| PDM3 |  | 6.804% | 22.19% | 9.36% | -5.545% | 2.254% | 1.388% | -8.144% | 5.721% |  | -8.144% |  | 4.793% | 2.254% | 95.84% | 3.988% | 2.254% |  |  | -9.011%
| Z-Score (Absolute) |  | 0.13 | 1.13 | 0.3 | 0.67 | 0.17 | 0.22 | 0.84 | 0.06 | 27.08 | 0.84 | 27.08 | 0 | 0.17 | 5.94 | 0.05 | 0.17 | 1.64 | 1.64 | 0.9
| Analytical results for Al in OREAS 290 (Certified Value 2.59 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g |  | 2
| 1 | 1 | 2.0055 | 2.32 | 2.6 | 2.51 | 2.5 | 2.55 | 2.46 | 2.33 | 2.31 | 2.47 | 3.051 | 2.539 | 3.13 | 2.3938 | 2.54 | 1.31 | 2.925 | 2.26 |  | 1
| 1 | 2 | 2.079 | 2.38 | 2.47 | 2.54 | 2.52 | 2.5 | 2.43 | 2.42 | 2.27 | 2.49 | 3.1559 | 2.547 | 3.1 | 3.0246 | 2.58 | 1.19 | 2.94 | 2.31 |  | 3
| 1 | 3 | 2.3205 | 2.41 | 2.59 | 2.51 | 2.53 | 2.58 | 2.48 | 2.38 | 2.28 | 2.52 | 3.0876 | 2.554 | 3.07 | 2.6768 | 2.88 | 1.79 | 2.95 | 2.42 |  | 16
| 1 | 4 | 2.079 | 2.37 | 2.62 | 2.56 | 2.54 | 2.68 | 2.52 | 2.41 | 2.29 | 2.58 | 3.0518 | 2.567 | 3.11 | 3.2739 | 2.71 | 1.84 | 2.95 | 2.72 |  | 2.5919
| 1 | 5 | 2.1 | 2.36 | 2.69 | 2.57 | 2.49 | 2.59 | 2.57 | 2.39 | 2.29 | 2.65 | 3.1005 | 2.544 | 3.15 | 2.5644 | 2.71 | 1.7 | 2.915 | 2.82 |  | 75
| 1 | 6 | 2.0475 | 2.42 | 2.68 | 2.48 | 2.53 | 2.53 | 2.53 | 2.38 | 2.3 | 2.65 | 3.1505 | 2.53 | 3.1 | 2.7856 | 2.75 | 1.51 | 2.905 | 2.44
| Mean |  | 2.1052 | 2.3767 | 2.6083 | 2.5283 | 2.5183 | 2.5717 | 2.4983 | 2.385 | 2.29 | 2.56 | 3.0996 | 2.5468 | 3.11 | 2.7865 | 2.695 | 1.5567 | 2.9308 | 2.495
| Median |  | 2.079 | 2.375 | 2.61 | 2.525 | 2.525 | 2.565 | 2.5 | 2.385 | 2.29 | 2.55 | 3.0941 | 2.5455 | 3.105 | 2.7312 | 2.71 | 1.605 | 2.9325 | 2.43
| Std Dev. |  | 0.11047 | 0.036148 | 0.079352 | 0.034303 | 0.019408 | 0.062423 | 0.051153 | 0.031464 | 0.014142 | 0.078994 | 0.045934 | 0.012734 | 0.027568 | 0.31932 | 0.12243 | 0.26561 | 0.018819 | 0.22554
| Rel.Std.Dev. |  | 5.248% | 1.521% | 3.042% | 1.357% | 0.7707% | 2.427% | 2.047% | 1.319% | 0.6176% | 3.086% | 1.482% | 0.5% | 0.8864% | 11.46% | 4.543% | 17.06% | 0.6421% | 9.04%
| PDM3 |  | -18.78% | -8.305% | 0.6333% | -2.453% | -2.839% | -0.7813% | -3.611% | -7.983% | -11.65% | -1.231% | 19.59% | -1.739% | 19.99% | 7.508% | 3.977% | -39.94% | 13.08% | -3.739%
| Z-Score (Absolute) |  | 1.88 | 0.7 | 0.31 | 0.04 | 0.08 | 0.15 | 0.17 | 0.66 | 1.08 | 0.1 | 2.44 | 0.04 | 2.49 | 1.08 | 0.68 | 4.27 | 1.71 | 0.19
| Analytical results for As in OREAS 290 (Certified Value 307 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g |  | 0
| 1 | 1 | 318.9286 | 321.11 | 301.3 | 331 | 304 | 304 | 323 | 335 | 250 | 286 | 296.6774 | 317.25 | 285 | 226 | 148 | 275 | 304.405 | 299 |  | 1
| 1 | 2 | 311.6834 | 322.98 | 304.8 | 312 | 300 | 304 | 317 | 320 | 267 | 292 | 300.6173 | 317.24 | 302 | 287 | 158 | 277 | 306.0221 | 335 |  | 4
| 1 | 3 | 299.388 | 307.38 | 303.1 | 312 | 308 | 303 | 317 | 331 | 248 | 286 | 294.0685 | 325.86 | 290 | 251 | 192 | 300 | 304.0625 | 310 |  | 16
| 1 | 4 | 310.4253 | 332.73 | 312.1 | 324 | 313 | 312 | 324 | 326 | 247 | 293 | 295.8268 | 325.3 | 296 | 304 | 156 | 324 | 301.06 | 310 |  | 306.7984
| 1 | 5 | 321.5909 | 325.13 | 304.8 | 327 | 309 | 302 | 323 | 324 | 258 | 291 | 292.7962 | 326.69 | 307 | 238 | 151 | 309 | 304.4833 | 322 |  | 76
| 1 | 6 | 310.758 | 318.41 | 299.9 | 319 | 305 | 304 | 316 | 334 | 244 | 290 | 300.9267 | 327.32 | 281 | 263 | 144 | 283 | 302.03 | 300
| Mean |  | 312.129 | 321.29 | 304.3333 | 320.8333 | 306.5 | 304.8333 | 320 | 328.3333 | 252.3333 | 289.6667 | 296.8188 | 323.2767 | 293.5 | 261.5 | 158.1667 | 294.6667 | 303.6771 | 312.6667
| Median |  | 311.2207 | 322.045 | 303.95 | 321.5 | 306.5 | 304 | 320 | 328.5 | 249 | 290.5 | 296.2521 | 325.58 | 293 | 257 | 153.5 | 291.5 | 304.2337 | 310
| Std Dev. |  | 7.7856 | 8.3701 | 4.2693 | 7.8846 | 4.5056 | 3.6009 | 3.6878 | 5.9554 | 8.5946 | 3.0111 | 3.3482 | 4.7229 | 10.015 | 29.6024 | 17.3484 | 19.6435 | 1.8105 | 13.7647
| Rel.Std.Dev. |  | 2.494% | 2.605% | 1.403% | 2.458% | 1.47% | 1.181% | 1.152% | 1.814% | 3.406% | 1.04% | 1.128% | 1.461% | 3.412% | 11.32% | 10.97% | 6.666% | 0.5962% | 4.402%
| PDM3 |  | 1.737% | 4.723% | -0.8035% | 4.575% | -0.09727% | -0.6405% | 4.303% | 7.019% | -17.75% | -5.584% | -3.253% | 5.371% | -4.335% | -14.76% | -48.45% | -3.954% | -1.017% | 1.913%
| Z-Score (Absolute) |  | 0.39 | 0.87 | 0.01 | 0.84 | 0.1 | 0.01 | 0.8 | 1.23 | 2.71 | 0.77 | 0.4 | 0.97 | 0.57 | 2.23 | 7.59 | 0.51 | 0.05 | 0.42
| Analytical results for B in OREAS 290 (Indicative Value 17.3 ppm) | [20 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | [6 empty] | 1
| 1 | 1 | 21 | < 10 | < 10 | < 10 | < 10 | < 10 | 17 | < 10 | < 5 | < 10 | 6 | 20.8 | 33.775 | [6 empty] | 1
| 1 | 2 | 22 | < 10 | < 10 | < 10 | < 10 | < 10 | 16 | < 10 | < 5 | < 10 | 6 | 21.2 | 33.3 | [6 empty] | 17
| 1 | 3 | 20 | < 10 | < 10 | < 10 | < 10 | < 10 | 16 | < 10 | < 5 | < 10 | 7 | 23.4 | 33.595 | [6 empty] | 16
| 1 | 4 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | 11 | < 10 | < 5 | < 10 | 7 | 23.9 | 32.905 | [6 empty] | 17.324
| 1 | 5 | 21 | < 10 | < 10 | < 10 | < 10 | < 10 | 14 | < 10 | < 5 | < 10 | 6 | 19.1 | 33.475 | [6 empty] | 23
| 1 | 6 | 22 | < 10 | 10 | < 10 | < 10 | < 10 | 16 | < 10 | < 5 | < 10 | 6 | 19.6 | 33.615
| Mean |  | 21.2 |  | 10 |  |  |  | 15 |  |  |  | 6.3333 | 21.3333 | 33.4442
| Median |  | 21 |  | 10 |  |  |  | 16 |  |  |  | 6 | 21 | 33.535
| Std Dev. |  | 0.83666 | [5 empty] | 2.1909 |  |  |  | 0.5164 | 1.9572 | 0.30784
| Rel.Std.Dev. |  | 3.947% | [5 empty] | 14.61% |  |  |  | 8.154% | 9.174% | 0.9205%
| PDM3 |  | 22.37% |  | -42.28% |  |  |  | -13.42% |  |  |  | -63.44% | 23.14% | 93.05%
| Z-Score (Absolute) |  | 11.6 | 0 | 0.67 | 0 | 0 | 0 | 8.09 | 0 | 2.02 | 0 | 1.08 | 13.22 | 23.02
| Analytical results for Ba in OREAS 290 (Certified Value 117 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 0
| 1 | 1 | 108.15 | 119.35 | 109.6 | 120 | 120 | 120 | 120 | 115 | 102 | 115 | 115 | 143.009 | 120.84 | 112 | 155 | 25 | 93 | 148.423 | 150 | 1
| 1 | 2 | 110.25 | 121.39 | 112.1 | 120 | 120 | 120 | 120 | 114 | 106 | 115 | 117 | 149.2634 | 118.74 | 121 | 143 | 43.4 | 87 | 148.426 | 137 | 5
| 1 | 3 | 116.55 | 122.01 | 112.2 | 120 | 120 | 120 | 120 | 116 | 104 | 111 | 119 | 147.2906 | 120.7 | 114 | 128 | 27.3 | 110 | 145.789 | 135 | 16
| 1 | 4 | 110.25 | 120.12 | 112.5 | 130 | 120 | 120 | 130 | 119 | 105 | 116 | 121 | 142.1543 | 123.43 | 121 | 155 | 26.5 | 110 | 146.035 | 144 | 116.6471
| 1 | 5 | 111.3 | 122.79 | 111.9 | 130 | 120 | 120 | 120 | 120 | 105 | 114 | 123 | 147.9117 | 123.83 | 122 | 125 | 27.9 | 110 | 149.635 | 155 | 77
| 1 | 6 | 110.25 | 124.69 | 113.5 | 130 | 120 | 120 | 120 | 118 | 104 | 118 | 125 | 151.4784 | 121.93 | 116 | 136 | 28.2 | 99 | 148.63 | 149
| Mean |  | 111.125 | 121.725 | 111.9667 | 125 | 120 | 120 | 121.6667 | 117 | 104.3333 | 114.8333 | 120 | 146.8512 | 121.5783 | 117.6667 | 140.3333 | 29.7167 | 101.5 | 147.823 | 145
| Median |  | 110.25 | 121.7 | 112.15 | 125 | 120 | 120 | 120 | 117 | 104.5 | 115 | 120 | 147.6011 | 121.385 | 118.5 | 139.5 | 27.6 | 104.5 | 148.4245 | 146.5
| Std Dev. |  | 2.8499 | 1.915 | 1.2894 | 5.4772 | 0 | 0 | 4.0825 | 2.3664 | 1.3663 | 2.3166 | 3.7417 | 3.6158 | 1.8971 | 4.2269 | 12.9872 | 6.801 | 10.0548 | 1.5487 | 7.823
| Rel.Std.Dev. |  | 2.565% | 1.573% | 1.152% | 4.382% | 0% | 0% | 3.355% | 2.023% | 1.31% | 2.017% | 3.118% | 2.462% | 1.56% | 3.592% | 9.255% | 22.89% | 9.906% | 1.048% | 5.395%
| PDM3 |  | -4.734% | 4.353% | -4.013% | 7.161% | 2.874% | 2.874% | 4.303% | 0.3025% | -10.56% | -1.555% | 2.874% | 25.89% | 4.227% | 0.874% | 20.31% | -74.52% | -12.99% | 26.73% | 24.31%
| Z-Score (Absolute) |  | 1.16 | 0.23 | 1.05 | 0.65 | 0 | 0 | 0.22 | 0.39 | 2.04 | 0.67 | 0 | 3.5 | 0.21 | 0.3 | 2.65 | 11.78 | 2.41 | 3.63 | 3.26
| Analytical results for Be in OREAS 290 (Certified Value 1.26 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.1g |  |  |  | 2
| 1 | 1 | 1.09 | 1.18 | 1.28 | 1.13 | 1.1 | 1.4 | 1.1 | 1.4 | 1.4 | 1.4961 | 1.38 | 1.27 | 1.3 | < 2 | < 5 | 0.81 |  |  |  | 1
| 1 | 2 | 1.11 | 1.12 | 1.3 | 1.25 | 1.05 | 1.3 | 1.1 | 1.4 | 1.4 | 1.56 | 1.38 | 1.31 | 1.5 | < 2 | < 5 | 0.91 |  |  |  | 6
| 1 | 3 | 1.06 | 1.26 | 1.27 | 1.33 | 1.07 | 1.1 | 1.1 | 1.3 | 1.4 | 1.5063 | 1.39 | 1.24 | 1.4 | < 2 | < 5 | 0.96 |  |  |  | 16
| 1 | 4 | 1.05 | 1.18 | 1.25 | 1.27 | 0.99 | 1.2 | 1.1 | 1.3 | 1.4 | 1.4965 | 1.38 | 1.29 | 1.4 | < 2 | < 5 | 0.92 |  |  |  | 1.2586
| 1 | 5 | 1.11 | 1.28 | 1.33 | 1.24 | 1 | 1.5 | 1.1 | 1.4 | 1.5 | 1.5767 | 1.38 | 1.31 | 1.4 | < 2 | < 5 | 0.86 |  |  |  | 78
| 1 | 6 | 1.12 | 1.44 | 1.29 | 1.28 | 1.02 | 1.4 | 1.1 | 1.5 | 1.5 | 1.5447 | 1.32 | 1.22 | 1.4 | < 2 | < 5 | 0.9
| Mean |  | 1.09 | 1.2433 | 1.2867 | 1.25 | 1.0383 | 1.3167 | 1.1 | 1.3833 | 1.4333 | 1.5301 | 1.3717 | 1.2733 | 1.4 |  |  | 0.89333
| Median |  | 1.1 | 1.22 | 1.285 | 1.26 | 1.035 | 1.35 | 1.1 | 1.4 | 1.4 | 1.5255 | 1.38 | 1.28 | 1.4 |  |  | 0.905
| Std Dev. |  | 0.028983 | 0.11272 | 0.027325 | 0.066633 | 0.042622 | 0.1472 | 2.4324e-16 | 0.075277 | 0.05164 | 0.035014 | 0.025626 | 0.037238 | 0.063246 |  |  | 0.052026
| Rel.Std.Dev. |  | 2.659% | 9.066% | 2.124% | 5.331% | 4.105% | 11.18% | 2.211e-14% | 5.442% | 3.603% | 2.288% | 1.868% | 2.924% | 4.518% |  |  | 5.824%
| PDM3 |  | -13.4% | -1.213% | 2.23% | -0.6832% | -17.5% | 4.614% | -12.6% | 9.911% | 13.88% | 21.57% | 8.984% | 1.171% | 11.23% |  |  | -29.02%
| Z-Score (Absolute) |  | 0.94 | 0.18 | 0.03 | 0.15 | 1.19 | 0.18 | 0.89 | 0.51 | 0.76 | 1.23 | 0.45 | 0.03 | 0.59 | 1.38 | 6.02 | 1.91
| Analytical results for Bi in OREAS 290 (Certified Value 0.36 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 2
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | < 3 | < 10 | 0.28 | 0.35 | 0.35 | 0.34 | 0.41 | 0.38 | < 10 | 0.34 | < 5 | 0.3801 | 0.35 | 0.4 | < 2 | 0.37 | < 2 | < 5 | 0.38 | 1
| 1 | 2 | < 3 | < 10 | 0.32 | 0.37 | 0.38 | 0.36 | 0.41 | 0.38 | < 10 | 0.35 | < 5 | 0.38185 | 0.35 | 0.38 | < 2 | 0.32 | < 2 | < 5 | 0.33 | 7
| 1 | 3 | < 3 | < 10 | 0.32 | 0.33 | 0.36 | 0.4 | 0.43 | 0.41 | < 10 | 0.32 | < 5 | 0.35483 | 0.35 | 0.37 | < 2 | 0.38 | < 2 | < 5 | 0.35 | 16
| 1 | 4 | < 3 | < 10 | 0.29 | 0.32 | 0.36 | 0.36 | 0.38 | 0.4 | < 10 | 0.32 | < 5 | 0.3662 | 0.35 | 0.37 | < 2 | 0.45 | < 2 | < 5 | 0.35 | 0.36432
| 1 | 5 | < 3 | < 10 | 0.28 | 0.34 | 0.38 | 0.37 | 0.33 | 0.37 | < 10 | 0.35 | < 5 | 0.36521 | 0.36 | 0.37 | < 2 | 0.38 | < 2 | < 5 | 0.39 | 79
| 1 | 6 | < 3 | < 10 | 0.28 | 0.37 | 0.4 | 0.35 | 0.36 | 0.37 | < 10 | 0.35 | < 5 | 0.36123 | 0.35 | 0.36 | < 2 | 0.38 | < 2 | < 5 | 0.36
| Mean |  |  |  | 0.295 | 0.34667 | 0.37167 | 0.36333 | 0.38667 | 0.385 |  | 0.33833 |  | 0.36823 | 0.35167 | 0.375 |  | 0.38 |  |  | 0.36
| Median |  |  |  | 0.285 | 0.345 | 0.37 | 0.36 | 0.395 | 0.38 |  | 0.345 |  | 0.3657 | 0.35 | 0.37 |  | 0.38 |  |  | 0.355
| Std Dev. |  |  |  | 0.019748 | 0.020656 | 0.018348 | 0.020656 | 0.037238 | 0.016432 |  | 0.01472 |  | 0.010658 | 0.0040825 | 0.013784 |  | 0.041473 |  |  | 0.021909
| Rel.Std.Dev. |  |  |  | 6.694% | 5.958% | 4.937% | 5.685% | 9.631% | 4.268% |  | 4.351% |  | 2.894% | 1.161% | 3.676% |  | 10.91% |  |  | 6.086%
| PDM3 |  |  |  | -19.03% | -4.847% | 2.015% | -0.272% | 6.133% | 5.675% |  | -7.134% |  | 1.073% | -3.474% | 2.93% |  | 4.303% |  |  | -1.187%
| Z-Score (Absolute) |  | 22.66 | 93.46 | 1.72 | 0.67 | 0.17 | 0.34 | 0.13 | 0.1 | 93.46 | 0.84 | 42.89 | 0.24 | 0.57 | 0.1 | 12.54 | 0 | 12.54 | 42.89 | 0.4
| Analytical results for Ca in OREAS 290 (Certified Value 0.312 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.3234 | 0.3445 | 0.29 | 0.32 | 0.31 | 0.31 | 0.3 | 0.31 | 0.28 | 0.31 | 0.29 | 0.31276 | 0.325 | 0.33 | 0.25 | 0.33 | 0.312 | 0.3038 | 0.266 | 1
| 1 | 2 | 0.3256 | 0.34125 | 0.3 | 0.31 | 0.31 | 0.31 | 0.3 | 0.31 | 0.28 | 0.3 | 0.29 | 0.31702 | 0.331 | 0.33 | 0.32 | 0.33 | 0.301 | 0.3136 | 0.293 | 8
| 1 | 3 | 0.3201 | 0.33587 | 0.3 | 0.31 | 0.31 | 0.32 | 0.31 | 0.32 | 0.28 | 0.3 | 0.3 | 0.31216 | 0.326 | 0.33 | 0.28 | 0.34 | 0.322 | 0.2989 | 0.283 | 16
| 1 | 4 | 0.3212 | 0.33616 | 0.3 | 0.32 | 0.32 | 0.32 | 0.33 | 0.32 | 0.28 | 0.3 | 0.29 | 0.31127 | 0.326 | 0.33 | 0.34 | 0.34 | 0.323 | 0.2989 | 0.308 | 0.31211
| 1 | 5 | 0.3289 | 0.34009 | 0.31 | 0.32 | 0.32 | 0.31 | 0.31 | 0.32 | 0.28 | 0.31 | 0.3 | 0.3115 | 0.328 | 0.34 | 0.26 | 0.33 | 0.321 | 0.3038 | 0.3017 | 80
| 1 | 6 | 0.3267 | 0.33984 | 0.3 | 0.32 | 0.31 | 0.31 | 0.3 | 0.32 | 0.28 | 0.31 | 0.3 | 0.3169 | 0.323 | 0.33 | 0.29 | 0.34 | 0.305 | 0.2989 | 0.325
| Mean |  | 0.32432 | 0.33962 | 0.3 | 0.31667 | 0.31333 | 0.31333 | 0.30833 | 0.31667 | 0.28 | 0.305 | 0.295 | 0.3136 | 0.3265 | 0.33167 | 0.29 | 0.335 | 0.314 | 0.30298 | 0.29612
| Median |  | 0.3245 | 0.33997 | 0.3 | 0.32 | 0.31 | 0.31 | 0.305 | 0.32 | 0.28 | 0.305 | 0.295 | 0.31246 | 0.326 | 0.33 | 0.285 | 0.335 | 0.3165 | 0.30135 | 0.29735
| Std Dev. |  | 0.0033666 | 0.0032498 | 0.0063246 | 0.005164 | 0.005164 | 0.005164 | 0.01169 | 0.005164 | 0 | 0.0054772 | 0.0054772 | 0.0026515 | 0.0027386 | 0.0040825 | 0.034641 | 0.0054772 | 0.0094657 | 0.0057283 | 0.020471
| Rel.Std.Dev. |  | 1.038% | 0.9569% | 2.108% | 1.631% | 1.648% | 1.648% | 3.791% | 1.631% | 0% | 1.796% | 1.857% | 0.8455% | 0.8388% | 1.231% | 11.95% | 1.635% | 3.015% | 1.891% | 6.913%
| PDM3 |  | 3.91% | 8.813% | -3.881% | 1.459% | 0.3911% | 0.3911% | -1.211% | 1.459% | -10.29% | -2.279% | -5.483% | 0.4771% | 4.61% | 6.265% | -7.085% | 7.333% | 0.6047% | -2.925% | -5.125%
| Z-Score (Absolute) |  | 0.67 | 1.61 | 0.82 | 0.2 | 0 | 0 | 0.31 | 0.2 | 2.05 | 0.51 | 1.13 | 0.02 | 0.81 | 1.13 | 1.43 | 1.33 | 0.04 | 0.64 | 1.06
| Analytical results for Cd in OREAS 290 (Certified Value 0.05 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | < 0.1 | 1.99 | 0.06 | 0.04 | 0.04 | 0.04 | 0.06 | 0.04 | < 1 | 0.07 | < 1 | < 0.05 | 0.05 | 0.07 | < 0.5 | 0.02 | < 2 | < 0.5 | 0.05 | 1
| 1 | 2 | < 0.1 | 2.31 | 0.08 | 0.04 | 0.04 | 0.05 | 0.04 | 0.06 | < 1 | 0.05 | < 1 | < 0.05 | 0.049 | 0.07 | < 0.5 | 0.03 | < 2 | < 0.5 | 0.06 | 9
| 1 | 3 | < 0.1 | 2.11 | 0.06 | 0.04 | 0.04 | 0.04 | 0.06 | 0.04 | < 1 | 0.05 | < 1 | < 0.05 | 0.047 | 0.07 | < 0.5 | 0.08 | < 2 | < 0.5 | 0.05 | 16
| 1 | 4 | < 0.1 | 2.25 | 0.06 | 0.05 | 0.05 | 0.05 | 0.04 | 0.05 | < 1 | 0.05 | < 1 | 0.054125 | 0.052 | 0.07 | < 0.5 | 0.04 | < 2 | < 0.5 | 0.05 | 0.050191
| 1 | 5 | < 0.1 | 1.72 | 0.06 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | < 1 | 0.05 | < 1 | 0.052542 | 0.051 | 0.07 | < 0.5 | 0.06 | < 2 | < 0.5 | 0.05 | 81
| 1 | 6 | < 0.1 | 1.72 | 0.07 | 0.05 | 0.04 | 0.04 | 0.04 | 0.03 | < 1 | 0.05 | < 1 | 0.050195 | 0.051 | 0.06 | < 0.5 | 0.02 | < 2 | < 0.5 | 0.04
| Mean |  |  | 2.0167 | 0.065 | 0.043333 | 0.041667 | 0.043333 | 0.048333 | 0.045 |  | 0.053333 |  | 0.052287 | 0.05 | 0.068333 |  | 0.041667 |  |  | 0.05
| Median |  |  | 2.05 | 0.06 | 0.04 | 0.04 | 0.04 | 0.045 | 0.045 |  | 0.05 |  | 0.052542 | 0.0505 | 0.07 |  | 0.035 |  |  | 0.05
| Std Dev. |  |  | 0.25532 | 0.0083666 | 0.005164 | 0.0040825 | 0.005164 | 0.0098319 | 0.010488 |  | 0.008165 |  | 0.0019776 | 0.0017889 | 0.0040825 |  | 0.024014 |  |  | 0.0063246
| Rel.Std.Dev. |  |  | 12.66% | 12.87% | 11.92% | 9.798% | 11.92% | 20.34% | 23.31% |  | 15.31% |  | 3.782% | 3.578% | 5.974% |  | 57.63% |  |  | 12.65%
| PDM3 |  |  | 3918% | 29.51% | -13.66% | -16.98% | -13.66% | -3.7% | -10.34% |  | 6.262% |  | 4.177% | -0.3798% | 36.15% |  | -16.98% |  |  | -0.3798%
| Z-Score (Absolute) |  | 0 | 159.14 | 1.21 | 0.54 | 0.67 | 0.54 | 0.13 | 0.4 | 36.41 | 0.27 | 36.41 | 0.92 | 0 | 1.48 | 16.18 | 0.67 | 76.87 | 16.18 | 0
| Analytical results for Ce in OREAS 290 (Certified Value 50 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.1g | [7 empty] | 0
| 1 | 1 | 47.9 | 42.7 | 44.6 | 45.3 | 39.27 | 36.62 | 63.0627 | 47.716 | 66.2 | 55 | 42.6 | 58.7 | [7 empty] | 1
| 1 | 2 | 43.7 | 44.2 | 46.4 | 41.7 | 38.47 | 37.72 | 67.4724 | 47.101 | 65.5 | 68 | 42.6 | 59 | [7 empty] | 38
| 1 | 3 | 51.1 | 45.1 | 48.2 | 44.3 | 40.35 | 36.27 | 64.3297 | 45.756 | 63.1 | 62 | 42.4 | 62.5 | [7 empty] | 16
| 1 | 4 | 48.1 | 44.5 | 45.9 | 39 | 42.42 | 36.31 | 62.8578 | 47.515 | 64.6 | 74 | 42.2 | 60.6 | [7 empty] | 50.1463
| 1 | 5 | 51.8 | 45.2 | 46 | 38 | 43.32 | 37.36 | 65.8637 | 48.079 | 65.3 | 59 | 42.6 | 60.3 | [7 empty] | 82
| 1 | 6 | 59.3 | 42.4 | 45.7 | 40.5 | 40.68 | 37.11 | 66.9232 | 47.927 | 63.8 | 63 | 42 | 54
| Mean |  | 50.3167 | 44.0167 | 46.1333 | 41.4667 | 40.7517 | 36.8983 | 65.0849 | 47.349 | 64.75 | 63.5 | 42.4 | 59.1833
| Median |  | 49.6 | 44.35 | 45.95 | 41.1 | 40.515 | 36.865 | 65.0967 | 47.6155 | 64.95 | 62.5 | 42.5 | 59.65
| Std Dev. |  | 5.2545 | 1.199 | 1.1793 | 2.8918 | 1.8411 | 0.59186 | 1.965 | 0.85187 | 1.1502 | 6.7157 | 0.25298 | 2.8771
| Rel.Std.Dev. |  | 10.44% | 2.724% | 2.556% | 6.974% | 4.518% | 1.604% | 3.019% | 1.799% | 1.776% | 10.58% | 0.5967% | 4.861%
| PDM3 |  | 0.3397% | -12.22% | -8.003% | -17.31% | -18.73% | -26.42% | 29.79% | -5.578% | 29.12% | 26.63% | -15.45% | 18.02%
| Z-Score (Absolute) |  | 0.43 | 0.33 | 0.07 | 0.63 | 0.72 | 1.18 | 2.2 | 0.07 | 2.16 | 2.01 | 0.52 | 1.49
| Analytical results for Co in OREAS 290 (Certified Value 13.4 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 2
| 1 | 1 | 13.86 | 16.67 | 12.8 | 13.1 | 13.4 | 12.5 | 15 | 13.5 | 11 | 12.4 | 13 | 13.4384 | 13.53 | 12.6 | 11 | 14.6 | 8.5 | 12.9575 | 13.6 | 1
| 1 | 2 | 13.7445 | 17.52 | 13.5 | 12.7 | 12.9 | 13.7 | 15.2 | 14.1 | 11 | 12.7 | 13 | 13.5216 | 13.4 | 13.6 | 13 | 14.1 | 8.2 | 13.0155 | 14.8 | 10
| 1 | 3 | 13.9755 | 18.28 | 13.1 | 12.8 | 13.3 | 14.2 | 13.1 | 14.2 | 11 | 12.1 | 13 | 13.4003 | 13.45 | 12.9 | 12 | 14.6 | 9.8 | 12.5382 | 13.6 | 16
| 1 | 4 | 13.629 | 17.56 | 13.4 | 11.8 | 13.2 | 13.8 | 12 | 14.3 | 11 | 12 | 13 | 13.4497 | 13.67 | 13.6 | 14 | 14.7 | 9.8 | 12.646 | 13.8 | 13.4206
| 1 | 5 | 13.9755 | 18.39 | 14 | 12.4 | 13.7 | 13.7 | 12 | 14.1 | 11 | 12.4 | 13 | 13.5269 | 13.62 | 13.2 | 11 | 14 | 9.7 | 12.518 | 14.1 | 83
| 1 | 6 | 13.7445 | 19.31 | 13.8 | 14.2 | 13 | 13.6 | 12.8 | 14 | 11 | 12.5 | 13 | 13.4192 | 13.47 | 12.8 | 12 | 14.1 | 8.9 | 12.646 | 13.6
| Mean |  | 13.8215 | 17.955 | 13.4333 | 12.8333 | 13.25 | 13.5833 | 13.35 | 14.0333 | 11 | 12.35 | 13 | 13.4594 | 13.5233 | 13.1167 | 12.1667 | 14.35 | 9.15 | 12.7202 | 13.9167
| Median |  | 13.8023 | 17.92 | 13.45 | 12.75 | 13.25 | 13.7 | 12.95 | 14.1 | 11 | 12.4 | 13 | 13.4441 | 13.5 | 13.05 | 12 | 14.35 | 9.3 | 12.646 | 13.7
| Std Dev. |  | 0.13988 | 0.90855 | 0.44121 | 0.80166 | 0.2881 | 0.57067 | 1.4251 | 0.28048 | 0 | 0.25884 | 0 | 0.053047 | 0.10424 | 0.42151 | 1.169 | 0.31464 | 0.71204 | 0.21378 | 0.47504
| Rel.Std.Dev. |  | 1.012% | 5.06% | 3.284% | 6.247% | 2.174% | 4.201% | 10.68% | 1.999% | 0% | 2.096% | 0% | 0.3941% | 0.7708% | 3.214% | 9.609% | 2.193% | 7.782% | 1.681% | 3.413%
| PDM3 |  | 2.988% | 33.79% | 0.09523% | -4.376% | -1.271% | 1.213% | -0.5257% | 4.566% | -18.04% | -7.977% | -3.134% | 0.2892% | 0.7658% | -2.264% | -9.343% | 6.926% | -31.82% | -5.218% | 3.697%
| Z-Score (Absolute) |  | 0.6 | 7.19 | 0.02 | 0.98 | 0.31 | 0.22 | 0.15 | 0.94 | NA | 1.75 | NA | 0.02 | 0.12 | 0.53 | NA | 1.44 | 6.85 | 1.16 | 0.75
Results from laboratories 2.10, 2.12 and 2.16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 290 (Certified Value 98 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 1
| 1 | 1 | 97.5975 | 94.93 | 88.6 | 101 | 99 | 98 | 100 | 102 | 86 | 99 | 100 | 99.8831 | 110.2 | 107 | 81 | 93 | 67.2 | 95.1012 | 72 | 1
| 1 | 2 | 96.558 | 96.07 | 89.8 | 97 | 99 | 99 | 99 | 104 | 88 | 97 | 100 | 99.4918 | 112.1 | 106 | 99 | 95 | 62 | 95.438 | 79 | 11
| 1 | 3 | 102.9105 | 97.34 | 91.9 | 99 | 99 | 101 | 101 | 103 | 86 | 92 | 101 | 101.2207 | 108.9 | 105 | 89 | 99 | 67.6 | 94.2225 | 78 | 16
| 1 | 4 | 95.9805 | 94.66 | 92 | 100 | 101 | 100 | 104 | 104 | 88 | 96 | 103 | 101.4064 | 109.7 | 107 | 99 | 100 | 70.4 | 94.4095 | 75 | 98.3039
| 1 | 5 | 96.6735 | 95.55 | 93 | 101 | 101 | 98 | 102 | 105 | 87 | 96 | 105 | 99.3289 | 113.1 | 108 | 85 | 94 | 66.8 | 94.0573 | 78 | 84
| 1 | 6 | 94.71 | 96.6 | 91.6 | 101 | 98 | 100 | 101 | 104 | 86 | 97 | 102 | 99.1471 | 110.3 | 106 | 93 | 99 | 55.7 | 94.118 | 76
| Mean |  | 97.405 | 95.8583 | 91.15 | 99.8333 | 99.5 | 99.3333 | 101.1667 | 103.6667 | 86.8333 | 96.1667 | 101.8333 | 100.0797 | 110.7167 | 106.5 | 91 | 96.6667 | 64.95 | 94.5577 | 76.3333
| Median |  | 96.6158 | 95.81 | 91.75 | 100.5 | 99 | 99.5 | 101 | 104 | 86.5 | 96.5 | 101.5 | 99.6875 | 110.25 | 106.5 | 91 | 97 | 67 | 94.316 | 77
| Std Dev. |  | 2.8598 | 1.0181 | 1.627 | 1.6021 | 1.2247 | 1.2111 | 1.7224 | 1.0328 | 0.98319 | 2.3166 | 1.9408 | 0.98789 | 1.5728 | 1.0488 | 7.3756 | 3.0111 | 5.2816 | 0.57419 | 2.582
| Rel.Std.Dev. |  | 2.936% | 1.062% | 1.785% | 1.605% | 1.231% | 1.219% | 1.703% | 0.9963% | 1.132% | 2.409% | 1.906% | 0.9871% | 1.421% | 0.9848% | 8.105% | 3.115% | 8.132% | 0.6072% | 3.383%
| PDM3 |  | -0.9144% | -2.488% | -7.277% | 1.556% | 1.217% | 1.047% | 2.912% | 5.455% | -11.67% | -2.174% | 3.59% | 1.806% | 12.63% | 8.338% | -7.43% | -1.665% | -33.93% | -3.811% | -22.35%
| Z-Score (Absolute) |  | 0 | 0.28 | 1.12 | 0.44 | 0.38 | 0.35 | 0.67 | 1.12 | 1.9 | 0.22 | 0.79 | 0.48 | 2.39 | 1.63 | 1.15 | 0.13 | 5.82 | 0.51 | 3.78
| Analytical results for Cs in OREAS 290 (Certified Value 7.24 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.20 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.1g | [8 empty] | 3
| 1 | 1 | 7.12 | 7.32 | 6.88 | 7.22 | 6.92 | 6.9 | 7.4619 | 7.44 | 7.02 | 7.25 | 7.7 | [8 empty] | 1
| 1 | 2 | 6.85 | 7.1 | 7.32 | 6.54 | 6.84 | 6.88 | 7.5267 | 7.42 | 7.62 | 7.05 | 8.2 | [8 empty] | 41
| 1 | 3 | 7.35 | 7.24 | 7.48 | 6.58 | 6.97 | 6.7 | 7.5542 | 7.34 | 7.15 | 7.15 | 7.8 | [8 empty] | 16
| 1 | 4 | 6.97 | 7.3 | 7.17 | 4.56 | 7.07 | 6.83 | 7.4603 | 7.57 | 7.3 | 7.14 | 7.8 | [8 empty] | 7.2424
| 1 | 5 | 7.25 | 7.39 | 7.15 | 5.31 | 7.03 | 6.9 | 7.543 | 7.62 | 7.6 | 7.42 | 7.6 | [8 empty] | 85
| 1 | 6 | 8.25 | 7.04 | 7.02 | 6.33 | 6.9 | 6.89 | 7.4599 | 7.46 | 7.19 | 7.11 | 7.5
| Mean |  | 7.2983 | 7.2317 | 7.17 | 6.09 | 6.955 | 6.85 | 7.501 | 7.475 | 7.3133 | 7.1867 | 7.7667
| Median |  | 7.185 | 7.27 | 7.16 | 6.435 | 6.945 | 6.885 | 7.4943 | 7.45 | 7.245 | 7.145 | 7.75
| Std Dev. |  | 0.50026 | 0.13541 | 0.21241 | 0.97226 | 0.085499 | 0.077974 | 0.045002 | 0.10271 | 0.24671 | 0.13155 | 0.24221
| Rel.Std.Dev. |  | 6.854% | 1.872% | 2.963% | 15.96% | 1.229% | 1.138% | 0.6% | 1.374% | 3.373% | 1.831% | 3.119%
| PDM3 |  | 0.7723% | -0.1482% | -0.9996% | -15.91% | -3.968% | -5.418% | 3.57% | 3.212% | 0.9794% | -0.7695% | 7.239%
| Z-Score (Absolute) |  | 0.18 | 0 | 0.17 | 3.16 | 0.77 | 1.06 | 0.75 | 0.67 | 0.23 | 0.12 | 1.48
| Analytical results for Cu in OREAS 290 (Certified Value 31.9 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 2
| 1 | 1 | 34.4651 | 13.94 | 31.82 | 30.6 | 32.7 | 31.7 | 39.1 | 32.8 | 27 | 34.9 | 31.8 | 29.7533 | 33.35 | 32 | 25 | 33.6 | 30.6 | 31.785 | 34.3 | 1
| 1 | 2 | 31.8834 | 14.56 | 31.69 | 29.9 | 31.7 | 33.3 | 36.2 | 33 | 28 | 34.8 | 31.5 | 30.4167 | 32.82 | 31 | 32 | 34.3 | 29.7 | 31.84 | 33.8 | 12
| 1 | 3 | 33.7045 | 12.95 | 32.48 | 31.8 | 33.1 | 34 | 35.6 | 33.1 | 28 | 32.8 | 31.6 | 29.8511 | 33.68 | 29 | 28 | 35.3 | 31.5 | 31.72 | 30.2 | 16
| 1 | 4 | 32.7017 | 11.82 | 32.78 | 28.7 | 32.6 | 33.3 | 33.5 | 34.7 | 27 | 33.1 | 31.7 | 29.2731 | 34.23 | 28 | 35 | 33.9 | 31.8 | 31.6 | 33 | 31.8903
| 1 | 5 | 32.3488 | 10.51 | 32.58 | 30 | 33.1 | 32.8 | 30.2 | 33.3 | 28 | 33.6 | 31.1 | 30.9157 | 35.07 | 30 | 26 | 34.2 | 32 | 30.955 | 32.5 | 86
| 1 | 6 | 32.8774 | 11.84 | 32.26 | 34.2 | 32.1 | 32.4 | 33.6 | 33.2 | 27 | 34.6 | 32.9 | 30.9893 | 34.28 | 30 | 29 | 34.9 | 29.1 | 30.99 | 31.9
| Mean |  | 32.9968 | 12.6033 | 32.2683 | 30.8667 | 32.55 | 32.9167 | 34.7 | 33.35 | 27.5 | 33.9667 | 31.7667 | 30.1999 | 33.905 | 30 | 29.1667 | 34.3667 | 30.7833 | 31.4817 | 32.6167
| Median |  | 32.7895 | 12.395 | 32.37 | 30.3 | 32.65 | 33.05 | 34.6 | 33.15 | 27.5 | 34.1 | 31.65 | 30.1339 | 33.955 | 30 | 28.5 | 34.25 | 31.05 | 31.66 | 32.75
| Std Dev. |  | 0.93997 | 1.5042 | 0.43342 | 1.9201 | 0.55767 | 0.80353 | 3.0106 | 0.68337 | 0.54772 | 0.91797 | 0.60553 | 0.6874 | 0.79238 | 1.4142 | 3.7639 | 0.6314 | 1.189 | 0.40258 | 1.4662
| Rel.Std.Dev. |  | 2.849% | 11.94% | 1.343% | 6.221% | 1.713% | 2.441% | 8.676% | 2.049% | 1.992% | 2.703% | 1.906% | 2.276% | 2.337% | 4.714% | 12.9% | 1.837% | 3.862% | 1.279% | 4.495%
| PDM3 |  | 3.47% | -60.48% | 1.185% | -3.21% | 2.069% | 3.219% | 8.811% | 4.577% | -13.77% | 6.511% | -0.3876% | -5.301% | 6.318% | -5.927% | -8.541% | 7.765% | -3.471% | -1.281% | 2.278%
| Z-Score (Absolute) |  | 0.33 | 8.93 | 0 | 0.64 | 0.13 | 0.29 | 1.1 | 0.49 | 2.17 | 0.77 | 0.23 | 0.94 | 0.74 | 1.03 | 1.41 | 0.95 | 0.67 | 0.36 | 0.16
| Analytical results for Dy in OREAS 290 (Indicative Value 1.87 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 1.949 | 1.9 | [17 empty] | 1
| 1 | 2 | 1.946 | 1.8 | [17 empty] | 18
| 1 | 3 | 1.896 | 1.7 | [17 empty] | 16
| 1 | 4 | 1.948 | 1.8 | [17 empty] | 1.8704
| 1 | 5 | 1.946 | 1.8 | [17 empty] | 24
| 1 | 6 | 1.96 | 1.8
| Mean |  | 1.9408 | 1.8
| Median |  | 1.947 | 1.8
| Std Dev. |  | 0.022578 | 0.063246
| Rel.Std.Dev. |  | 1.163% | 3.514%
| PDM3 |  | 3.765% | -3.765%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Er in OREAS 290 (Indicative Value 0.87 ppm) | [20 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.883 | 0.9 | [17 empty] | 1
| 1 | 2 | 0.841 | 0.9 | [17 empty] | 19
| 1 | 3 | 0.853 | 0.9 | [17 empty] | 16
| 1 | 4 | 0.882 | 0.9 | [17 empty] | 0.87017
| 1 | 5 | 0.908 | 0.8 | [17 empty] | 25
| 1 | 6 | 0.875 | 0.8
| Mean |  | 0.87367 | 0.86667
| Median |  | 0.8785 | 0.9
| Std Dev. |  | 0.023796 | 0.05164
| Rel.Std.Dev. |  | 2.724% | 5.958%
| PDM3 |  | 0.4022% | -0.4022%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 290 (Indicative Value 0.53 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.572 | 0.5 | [17 empty] | 1
| 1 | 2 | 0.512 | 0.5 | [17 empty] | 20
| 1 | 3 | 0.568 | 0.5 | [17 empty] | 16
| 1 | 4 | 0.559 | 0.5 | [17 empty] | 0.52942
| 1 | 5 | 0.571 | 0.5 | [17 empty] | 26
| 1 | 6 | 0.571 | 0.5
| Mean |  | 0.55883 | 0.5
| Median |  | 0.5695 | 0.5
| Std Dev. |  | 0.023439 | 0
| Rel.Std.Dev. |  | 4.194% | 0%
| PDM3 |  | 5.556% | -5.556%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Fe in OREAS 290 (Certified Value 3.38 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 2
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 3.0996 | 3.2768 | 3.13 | 3.41 | 3.35 | 3.25 | 3.31 | 3.41 | 2.95 | 3.3 | 3.34 | 3.5296 | 3.562 | 3.68 | 3.33 | 3.5 | 1.9 | 3.4 | 3.37 | 1
| 1 | 2 | 3.1428 | 3.2797 | 3.16 | 3.31 | 3.36 | 3.25 | 3.32 | 3.43 | 3.04 | 3.32 | 3.33 | 3.553 | 3.465 | 3.64 | 3.72 | 3.47 | 1.79 | 3.41 | 3.59 | 13
| 1 | 3 | 3.3048 | 3.2604 | 3.2 | 3.36 | 3.34 | 3.29 | 3.38 | 3.45 | 2.98 | 3.35 | 3.37 | 3.5332 | 3.541 | 3.69 | 3.26 | 3.64 | 2.3 | 3.365 | 3.56 | 16
| 1 | 4 | 3.1212 | 3.235 | 3.21 | 3.39 | 3.41 | 3.28 | 3.47 | 3.46 | 3 | 3.3 | 3.39 | 3.5195 | 3.583 | 3.75 | 4.04 | 3.58 | 2.35 | 3.36 | 3.86 | 3.3821
| 1 | 5 | 3.1644 | 3.3071 | 3.19 | 3.42 | 3.38 | 3.22 | 3.37 | 3.51 | 2.98 | 3.31 | 3.44 | 3.51 | 3.664 | 3.76 | 3.09 | 3.52 | 2.28 | 3.405 | 3.7825 | 87
| 1 | 6 | 3.1104 | 3.3271 | 3.2 | 3.41 | 3.31 | 3.26 | 3.33 | 3.49 | 2.96 | 3.32 | 3.42 | 3.5409 | 3.587 | 3.66 | 3.38 | 3.6 | 2.08 | 3.4 | 3.98
| Mean |  | 3.1572 | 3.281 | 3.1817 | 3.3833 | 3.3583 | 3.2583 | 3.3633 | 3.4583 | 2.985 | 3.3167 | 3.3817 | 3.531 | 3.567 | 3.6967 | 3.47 | 3.5517 | 2.1167 | 3.39 | 3.6904
| Median |  | 3.132 | 3.2782 | 3.195 | 3.4 | 3.355 | 3.255 | 3.35 | 3.455 | 2.98 | 3.315 | 3.38 | 3.5314 | 3.5725 | 3.685 | 3.355 | 3.55 | 2.18 | 3.4 | 3.6863
| Std Dev. |  | 0.075959 | 0.032759 | 0.030605 | 0.041793 | 0.034303 | 0.024833 | 0.059217 | 0.037103 | 0.032094 | 0.018619 | 0.043551 | 0.015219 | 0.065069 | 0.048442 | 0.34745 | 0.065243 | 0.23226 | 0.021679 | 0.22399
| Rel.Std.Dev. |  | 2.406% | 0.9984% | 0.9619% | 1.235% | 1.021% | 0.7621% | 1.761% | 1.073% | 1.075% | 0.5614% | 1.288% | 0.431% | 1.824% | 1.31% | 10.01% | 1.837% | 10.97% | 0.6395% | 6.069%
| PDM3 |  | -6.65% | -2.99% | -5.927% | 0.03593% | -0.7032% | -3.66% | -0.5554% | 2.253% | -11.74% | -1.935% | -0.01334% | 4.403% | 5.466% | 9.3% | 2.598% | 5.013% | -37.42% | 0.233% | 9.116%
| Z-Score (Absolute) |  | 1.23 | 0.55 | 1.09 | 0.01 | 0.13 | 0.67 | 0.1 | 0.42 | 2.17 | 0.36 | 0 | 0.82 | 1.01 | 1.72 | 0.48 | 0.93 | 6.92 | 0.05 | 1.69
| Analytical results for Ga in OREAS 290 (Certified Value 8.06 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [6 empty] | 3
| 1 | 1 | 7.6 | 7.83 | 8.41 | 6.94 | 7.75 | 7.6 | 7.3 | 8.7922 | 8.1 | 7.1 | 8.4 | 8.952 | 8.37 | [6 empty] | 1
| 1 | 2 | 7.7 | 7.57 | 8.37 | 7.67 | 7.39 | 7.6 | 7.4 | 9.1859 | 7.9 | 7.7 | 8.04 | 8.848 | 8.49 | [6 empty] | 46
| 1 | 3 | 8.4 | 7.34 | 8.45 | 7.75 | 7.81 | 7.8 | 7.1 | 8.8255 | 8.1 | 7.7 | 8.62 | 8.98 | 8.71 | [6 empty] | 16
| 1 | 4 | 8 | 7.15 | 8.46 | 7.62 | 7.48 | 7.8 | 7.1 | 8.441 | 8.3 | 7.8 | 8.36 | 8.78 | 8.64 | [6 empty] | 8.0593
| 1 | 5 | 8.3 | 7.34 | 8.62 | 7.57 | 6.83 | 8.1 | 7.6 | 8.5519 | 8.4 | 7.6 | 8.41 | 8.88 | 8.48 | [6 empty] | 88
| 1 | 6 | 8.4 | 8.3 | 8.26 | 7.34 | 7.47 | 7.8 | 7.7 | 9.0671 | 8.2 | 7.7 | 8.43 | 8.96 | 7.84
| Mean |  | 8.0667 | 7.5883 | 8.4283 | 7.4817 | 7.455 | 7.7833 | 7.3667 | 8.8106 | 8.1667 | 7.6 | 8.3767 | 8.9 | 8.4217
| Median |  | 8.15 | 7.455 | 8.43 | 7.595 | 7.475 | 7.8 | 7.35 | 8.8089 | 8.15 | 7.7 | 8.405 | 8.916 | 8.485
| Std Dev. |  | 0.3559 | 0.41969 | 0.11856 | 0.29929 | 0.34893 | 0.18348 | 0.25033 | 0.2866 | 0.17512 | 0.25298 | 0.18833 | 0.077728 | 0.3098
| Rel.Std.Dev. |  | 4.412% | 5.531% | 1.407% | 4% | 4.68% | 2.357% | 3.398% | 3.253% | 2.144% | 3.329% | 2.248% | 0.8733% | 3.679%
| PDM3 |  | 0.09201% | -5.843% | 4.58% | -7.167% | -7.498% | -3.424% | -8.594% | 9.323% | 1.333% | -5.698% | 3.939% | 10.43% | 4.497%
| Z-Score (Absolute) |  | 0 | 0.69 | 0.52 | 0.85 | 0.88 | 0.41 | 1.01 | 1.07 | 0.14 | 0.67 | 0.45 | 1.2 | 0.51
| Analytical results for Gd in OREAS 290 (Indicative Value 2.96 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 3.027 | 2.9 | [17 empty] | 1
| 1 | 2 | 2.895 | 2.9 | [17 empty] | 21
| 1 | 3 | 2.959 | 2.9 | [17 empty] | 16
| 1 | 4 | 3.101 | 2.9 | [17 empty] | 2.9587
| 1 | 5 | 3.073 | 2.9 | [17 empty] | 27
| 1 | 6 | 3.049 | 2.9
| Mean |  | 3.0173 | 2.9
| Median |  | 3.038 | 2.9
| Std Dev. |  | 0.076894 | 0
| Rel.Std.Dev. |  | 2.548% | 0%
| PDM3 |  | 1.983% | -1.983%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ge in OREAS 290 (Certified Value 0.097 ppm) | [20 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.17 | 2.20 | [10 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.1g | [10 empty] | 3
| 1 | 1 | 0.1 | 0.09 | 0.09 | 0.22 | 0.2 | 0.1 | < 0.05 | < 0.1 | 0.1 | [10 empty] | 1
| 1 | 2 | 0.1 | 0.09 | 0.09 | 0.23 | 0.2 | 0.1 | < 0.05 | < 0.1 | 0.09 | [10 empty] | 48
| 1 | 3 | 0.1 | 0.11 | 0.1 | 0.21 | 0.1 | 0.1 | < 0.05 | < 0.1 | 0.1 | [10 empty] | 16
| 1 | 4 | 0.09 | 0.09 | 0.09 | 0.21 | 0.1 | 0.1 | < 0.05 | < 0.1 | 0.11 | [10 empty] | 0.096667
| 1 | 5 | 0.1 | 0.09 | 0.1 | 0.22 | 0.1 | 0.1 | < 0.05 | < 0.1 | 0.09 | [10 empty] | 89
| 1 | 6 | 0.11 | 0.09 | 0.09 | 0.21 | 0.1 | 0.1 | < 0.05 | < 0.1 | 0.09
| Mean |  | 0.1 | 0.093333 | 0.093333 | 0.21667 | 0.13333 | 0.1 |  |  | 0.096667
| Median |  | 0.1 | 0.09 | 0.09 | 0.215 | 0.1 | 0.1 |  |  | 0.095
| Std Dev. |  | 0.0063246 | 0.008165 | 0.005164 | 0.008165 | 0.05164 | 1.5202e-17 |  |  | 0.008165
| Rel.Std.Dev. |  | 6.325% | 8.748% | 5.533% | 3.768% | 38.73% | 1.52e-14% |  |  | 8.447%
| PDM3 |  | 3.448% | -3.448% | -3.448% | 124.1% | 37.93% | 3.448% |  |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 24.28 | 7.42 | 0.67 | 14.5 | 9.44 | 0
| Analytical results for Hf in OREAS 290 (Certified Value 0.51 ppm) | [20 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.20 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.1g | [8 empty] | 2
| 1 | 1 | 0.46 | 0.48 | 0.48 | 0.49 | 0.5 | 0.52 | 0.86317 | 0.53 | 0.71 | < 0.1 | 0.62 | [8 empty] | 1
| 1 | 2 | 0.44 | 0.48 | 0.5 | 0.42 | 0.48 | 0.52 | 0.87615 | 0.57 | 0.78 | < 0.1 | 0.59 | [8 empty] | 32
| 1 | 3 | 0.46 | 0.51 | 0.53 | 0.48 | 0.55 | 0.49 | 0.88919 | 0.55 | 0.76 | < 0.1 | 0.76 | [8 empty] | 16
| 1 | 4 | 0.4 | 0.51 | 0.49 | 0.47 | 0.57 | 0.5 | 0.86391 | 0.58 | 0.73 | < 0.1 | 0.8 | [8 empty] | 0.51173
| 1 | 5 | 0.43 | 0.51 | 0.49 | 0.43 | 0.57 | 0.53 | 0.87147 | 0.54 | 0.74 | < 0.1 | 0.63 | [8 empty] | 90
| 1 | 6 | 0.5 | 0.49 | 0.48 | 0.41 | 0.52 | 0.52 | 0.89559 | 0.55 | 0.72 | < 0.1 | 0.61
| Mean |  | 0.44833 | 0.49667 | 0.495 | 0.45 | 0.53167 | 0.51333 | 0.87658 | 0.55333 | 0.74 |  | 0.66833
| Median |  | 0.45 | 0.5 | 0.49 | 0.45 | 0.535 | 0.52 | 0.87381 | 0.55 | 0.735 |  | 0.625
| Std Dev. |  | 0.033714 | 0.015055 | 0.018708 | 0.034059 | 0.037639 | 0.015055 | 0.013321 | 0.018619 | 0.026077 |  | 0.088412
| Rel.Std.Dev. |  | 7.52% | 3.031% | 3.779% | 7.569% | 7.079% | 2.933% | 1.52% | 3.365% | 3.524% |  | 13.23%
| PDM3 |  | -12.39% | -2.943% | -3.269% | -12.06% | 3.896% | 0.3135% | 71.3% | 8.13% | 44.61% |  | 30.6%
| Z-Score (Absolute) |  | 0.69 | 0.18 | 0.2 | 0.67 | 0.2 | 0 | 3.87 | 0.43 | 2.41 | 4.93 | 1.65
| Analytical results for Hg in OREAS 290 (Certified Value 0.031 ppm) | [20 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.20 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 0.5g | 0.1g | [6 empty] | 3
| 1 | 1 | 0.035 | 0.03 | 0.03 | 0.04 | 0.15 | 0.02 | < 5 | 0.06 | < 1 | < 0.05 | 0.03 | 0.04 | 0.03 | [6 empty] | 1
| 1 | 2 | 0.039 | 0.02 | 0.03 | 0.03 | 0.14 | 0.03 | < 5 | 0.05 | < 1 | < 0.05 | 0.03 | 0.03 | 0.03 | [6 empty] | 15
| 1 | 3 | 0.038 | 0.03 | 0.03 | 0.03 | 0.13 | < 0.01 | < 5 | 0.04 | < 1 | < 0.05 | 0.03 | 0.04 | 0.03 | [6 empty] | 16
| 1 | 4 | 0.036 | 0.02 | 0.03 | 0.03 | 0.11 | 0.01 | < 5 | 0.04 | < 1 | < 0.05 | 0.03 | 0.02 | 0.04 | [6 empty] | 0.030738
| 1 | 5 | 0.033 | 0.03 | 0.03 | 0.02 | 0.15 | 0.01 | < 5 | 0.04 | < 1 | < 0.05 | 0.03 | 0.02 | 0.04 | [6 empty] | 91
| 1 | 6 | 0.04 | 0.04 | 0.02 | 0.02 | 0.17 | 0.02 | < 5 | 0.04 | < 1 | < 0.05 | 0.03 | 0.03 | 0.03
| Mean |  | 0.036833 | 0.028333 | 0.028333 | 0.028333 | 0.14167 | 0.018 |  | 0.045 |  |  | 0.03 | 0.03 | 0.033333
| Median |  | 0.037 | 0.03 | 0.03 | 0.03 | 0.145 | 0.02 |  | 0.04 |  |  | 0.03 | 0.03 | 0.03
| Std Dev. |  | 0.0026394 | 0.0075277 | 0.0040825 | 0.0075277 | 0.020412 | 0.0083666 |  | 0.0083666 |  |  | 0 | 0.0089443 | 0.005164
| Rel.Std.Dev. |  | 7.166% | 26.57% | 14.41% | 26.57% | 14.41% | 46.48% |  | 18.59% |  |  | 0% | 29.81% | 15.49%
| PDM3 |  | 19.83% | -7.823% | -7.823% | -7.823% | 360.9% | -41.44% |  | 46.4% |  |  | -2.401% | -2.401% | 8.443%
| Z-Score (Absolute) |  | 0.92 | 0.22 | 0.22 | 0.22 | 15.06 | 1.91 | 333.11 | 2.02 | 63.39 | 0.67 | 0 | 0 | 0.45
| Analytical results for Ho in OREAS 290 (Indicative Value 0.32 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.346 | 0.3 | [17 empty] | 1
| 1 | 2 | 0.346 | 0.3 | [17 empty] | 17
| 1 | 3 | 0.332 | 0.3 | [17 empty] | 16
| 1 | 4 | 0.345 | 0.3 | [17 empty] | 0.32175
| 1 | 5 | 0.35 | 0.3 | [17 empty] | 23
| 1 | 6 | 0.342 | 0.3
| Mean |  | 0.3435 | 0.3
| Median |  | 0.3455 | 0.3
| Std Dev. |  | 0.0061887 | 0
| Rel.Std.Dev. |  | 1.802% | 0%
| PDM3 |  | 6.76% | -6.76%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 290 (Certified Value 0.034 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [7 empty] | 3
| 1 | 1 | 0.034 | 0.036 | 0.028 | 0.027 | 0.03 | 0.03 | < 0.05 | 0.03 | < 0.05 | 0.04 | < 5 | 0.04 | [7 empty] | 1
| 1 | 2 | 0.031 | 0.03 | 0.033 | 0.026 | 0.03 | 0.04 | < 0.05 | 0.03 | < 0.05 | 0.04 | < 5 | 0.04 | [7 empty] | 34
| 1 | 3 | 0.037 | 0.032 | 0.035 | 0.06 | 0.04 | 0.03 | < 0.05 | 0.04 | < 0.05 | 0.04 | < 5 | 0.03 | [7 empty] | 16
| 1 | 4 | 0.033 | 0.031 | 0.034 | 0.029 | 0.04 | 0.03 | < 0.05 | 0.04 | < 0.05 | 0.04 | < 5 | 0.03 | [7 empty] | 0.03423
| 1 | 5 | 0.035 | 0.034 | 0.032 | 0.04 | 0.03 | 0.04 | < 0.05 | 0.04 | < 0.05 | 0.03 | < 5 | 0.04 | [7 empty] | 92
| 1 | 6 | 0.039 | 0.033 | 0.029 | 0.03 | 0.03 | 0.04 | < 0.05 | 0.03 | < 0.05 | 0.04 | < 5 | 0.04
| Mean |  | 0.034833 | 0.032667 | 0.031833 | 0.035333 | 0.033333 | 0.035 |  | 0.035 |  | 0.038333 |  | 0.036667
| Median |  | 0.0345 | 0.0325 | 0.0325 | 0.0295 | 0.03 | 0.035 |  | 0.035 |  | 0.04 |  | 0.04
| Std Dev. |  | 0.0028577 | 0.0021602 | 0.0027869 | 0.013079 | 0.005164 | 0.0054772 |  | 0.0054772 |  | 0.0040825 |  | 0.005164
| Rel.Std.Dev. |  | 8.204% | 6.613% | 8.755% | 37.02% | 15.49% | 15.65% |  | 15.65% |  | 10.65% |  | 14.08%
| PDM3 |  | 1.764% | -4.566% | -7.001% | 3.224% | -2.618% | 2.251% |  | 2.251% |  | 11.99% |  | 7.12%
| Z-Score (Absolute) |  | 0.03 | 0.76 | 1.04 | 0.14 | 0.53 | 0.03 | 3.34 | 0.03 | 3.34 | 1.15 | 831.11 | 0.59
| Analytical results for K in OREAS 290 (Certified Value 0.858 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.8173 | 0.79451 | 0.82 | 0.85 | 0.85 | 0.84 | 0.87 | 0.86 | 0.83 | 0.86 | 0.85 | 1.002 | 0.917 | 0.885 | 0.9777 | 0.78 | 0.717 | 0.935 | 0.893 | 1
| 1 | 2 | 0.8382 | 0.81766 | 0.82 | 0.84 | 0.86 | 0.85 | 0.85 | 0.86 | 0.85 | 0.85 | 0.85 | 1.0448 | 0.852 | 0.877 | 0.9556 | 0.76 | 0.68 | 0.94 | 0.918 | 4
| 1 | 3 | 0.8591 | 0.81912 | 0.82 | 0.86 | 0.85 | 0.85 | 0.89 | 0.87 | 0.84 | 0.82 | 0.85 | 1.0165 | 0.89 | 0.869 | 0.97 | 0.81 | 0.745 | 0.915 | 0.94 | 16
| 1 | 4 | 0.8382 | 0.81949 | 0.83 | 0.86 | 0.87 | 0.85 | 0.89 | 0.88 | 0.85 | 0.85 | 0.87 | 0.99623 | 0.894 | 0.881 | 0.9621 | 0.79 | 0.753 | 0.91 | 1.01 | 0.85765
| 1 | 5 | 0.858 | 0.85022 | 0.82 | 0.87 | 0.87 | 0.84 | 0.87 | 0.89 | 0.85 | 0.85 | 0.89 | 1.0196 | 0.897 | 0.893 | 0.9639 | 0.84 | 0.748 | 0.92 | 1.08 | 93
| 1 | 6 | 0.8547 | 0.85525 | 0.84 | 0.87 | 0.85 | 0.86 | 0.87 | 0.88 | 0.84 | 0.86 | 0.89 | 1.0359 | 0.89 | 0.872 | 0.9432 | 0.77 | 0.718 | 0.93 | 0.989
| Mean |  | 0.84425 | 0.82604 | 0.825 | 0.85833 | 0.85833 | 0.84833 | 0.87333 | 0.87333 | 0.84333 | 0.84833 | 0.86667 | 1.0192 | 0.89 | 0.8795 | 0.96208 | 0.79167 | 0.72683 | 0.925 | 0.97167
| Median |  | 0.84645 | 0.81931 | 0.82 | 0.86 | 0.855 | 0.85 | 0.87 | 0.875 | 0.845 | 0.85 | 0.86 | 1.0181 | 0.892 | 0.879 | 0.963 | 0.785 | 0.7315 | 0.925 | 0.9645
| Std Dev. |  | 0.016238 | 0.022776 | 0.0083666 | 0.01169 | 0.0098319 | 0.0075277 | 0.015055 | 0.012111 | 0.008165 | 0.01472 | 0.019664 | 0.018797 | 0.021157 | 0.0088034 | 0.011904 | 0.029269 | 0.02768 | 0.011832 | 0.068669
| Rel.Std.Dev. |  | 1.923% | 2.757% | 1.014% | 1.362% | 1.145% | 0.8874% | 1.724% | 1.387% | 0.9682% | 1.735% | 2.269% | 1.844% | 2.377% | 1.001% | 1.237% | 3.697% | 3.808% | 1.279% | 7.067%
| PDM3 |  | -1.562% | -3.685% | -3.807% | 0.07986% | 0.07986% | -1.086% | 1.829% | 1.829% | -1.669% | -1.086% | 1.052% | 18.83% | 3.772% | 2.548% | 12.18% | -7.693% | -15.25% | 7.853% | 13.29%
| Z-Score (Absolute) |  | 0.45 | 1.03 | 1.06 | 0 | 0 | 0.32 | 0.48 | 0.48 | 0.48 | 0.32 | 0.27 | 5.12 | 1.01 | 0.67 | 3.31 | 2.12 | 4.19 | 2.12 | 3.61
| Analytical results for La in OREAS 290 (Certified Value 23.7 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.1g |  |  | 1
| 1 | 1 | 13.0615 | 18.2 | 24 | 22.7 | 22.7 | 23.8 | 20.7 | 18.8 | 19.8 | 24.8 | 32.1992 | 24.219 | 31.6 | 29 | 20.9 | 23.285 | 31.9 |  |  | 1
| 1 | 2 | 12.5515 | 18.7 | 22.1 | 23.1 | 24.5 | 21.6 | 20.3 | 20 | 19.9 | 25 | 34.1595 | 24.417 | 31.8 | 36 | 21 | 23.525 | 32.1 |  |  | 5
| 1 | 3 | 12.3272 | 19.1 | 25.9 | 23.5 | 25 | 23.4 | 21.2 | 19.1 | 18.2 | 24.8 | 32.4961 | 23.417 | 30.3 | 32 | 21.1 | 23.095 | 33.3 |  |  | 16
| 1 | 4 | 13.1122 | 18.6 | 24.3 | 23.4 | 23.6 | 19.6 | 22.4 | 20 | 18.4 | 26.2 | 32.0629 | 24.208 | 31.7 | 38 | 20.8 | 23.3167 | 33.1 |  |  | 23.7008
| 1 | 5 | 13.0306 | 18.1 | 26.2 | 23.9 | 23.4 | 20.7 | 22.6 | 20.4 | 19.6 | 28.7 | 33.0988 | 24.501 | 31.4 | 30 | 20.9 | 23.34 | 32.1 |  |  | 94
| 1 | 6 | 12.759 | 19.7 | 29.9 | 22.5 | 22.9 | 21.4 | 21.4 | 19.8 | 19.8 | 28.7 | 33.751 | 24.313 | 31.5 | 33 | 20.9 | 23.815 | 29.7
| Mean |  | 12.807 | 18.7333 | 25.4 | 23.1833 | 23.6833 | 21.75 | 21.4333 | 19.6833 | 19.2833 | 26.3667 | 32.9613 | 24.1792 | 31.3833 | 33 | 20.9333 | 23.3961 | 32.0333
| Median |  | 12.8948 | 18.65 | 25.1 | 23.25 | 23.5 | 21.5 | 21.3 | 19.9 | 19.7 | 25.6 | 32.7974 | 24.266 | 31.55 | 32.5 | 20.9 | 23.3283 | 32.1
| Std Dev. |  | 0.318 | 0.59554 | 2.6533 | 0.52313 | 0.90203 | 1.5997 | 0.9136 | 0.608 | 0.7705 | 1.8811 | 0.8582 | 0.39022 | 0.54924 | 3.4641 | 0.10328 | 0.2468 | 1.2817
| Rel.Std.Dev. |  | 2.483% | 3.179% | 10.45% | 2.256% | 3.809% | 7.355% | 4.263% | 3.089% | 3.996% | 7.135% | 2.604% | 1.614% | 1.75% | 10.5% | 0.4934% | 1.055% | 4.001%
| PDM3 |  | -45.96% | -20.96% | 7.169% | -2.183% | -0.07365% | -8.231% | -9.567% | -16.95% | -18.64% | 11.25% | 39.07% | 2.018% | 32.41% | 39.24% | -11.68% | -1.286% | 35.16%
| Z-Score (Absolute) |  | 2.4 | 1.06 | 0.45 | 0.05 | 0.07 | 0.37 | 0.45 | 0.84 | 0.93 | 0.67 | 2.17 | 0.18 | 1.81 | 2.18 | 0.56 | 0 | 1.96
| Analytical results for Li in OREAS 290 (Certified Value 37.6 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g |  |  |  | 1
| 1 | 1 | 37 | 37.1 | 32.5 | 36.2 | 40 | 26 | 33 | 39 | 43.1823 | 35.94 | 42.7 | 33 | 39.8 | 18.6 | 42.675 | 39.3 |  |  |  | 1
| 1 | 2 | 35.9 | 36.3 | 35.8 | 34.4 | 40 | 28 | 33 | 39 | 44.661 | 35.71 | 43.3 | 41 | 39.3 | 17.8 | 42.1617 | 40 |  |  |  | 6
| 1 | 3 | 36.7 | 36.5 | 37.2 | 34.6 | 41 | 27 | 31 | 39 | 44.5361 | 36.39 | 41.1 | 36 | 43.1 | 22.3 | 41.64 | 41.1 |  |  |  | 16
| 1 | 4 | 33.9 | 36.3 | 35.3 | 26.9 | 41 | 28 | 32 | 40 | 43.2074 | 36.45 | 42.7 | 44 | 42.1 | 22.7 | 41.53 | 42.3 |  |  |  | 37.5991
| 1 | 5 | 36.1 | 37.4 | 34.9 | 29.1 | 42 | 28 | 33 | 41 | 44.4066 | 37.02 | 42.2 | 34 | 39.9 | 21.1 | 42.105 | 40.2 |  |  |  | 95
| 1 | 6 | 41.1 | 35.6 | 34.8 | 32.8 | 42 | 28 | 33 | 41 | 45.3718 | 35.79 | 39.4 | 37 | 41.9 | 19.6 | 42.12 | 36.4
| Mean |  | 36.7833 | 36.5333 | 35.0833 | 32.3333 | 41 | 27.5 | 32.5 | 39.8333 | 44.2275 | 36.2167 | 41.9 | 37.5 | 41.0167 | 20.35 | 42.0386 | 39.8833
| Median |  | 36.4 | 36.4 | 35.1 | 33.6 | 41 | 28 | 33 | 39.5 | 44.4714 | 36.165 | 42.45 | 36.5 | 40.9 | 20.35 | 42.1125 | 40.1
| Std Dev. |  | 2.3769 | 0.64083 | 1.5381 | 3.5931 | 0.89443 | 0.83666 | 0.83666 | 0.98319 | 0.86693 | 0.49903 | 1.4297 | 4.2308 | 1.5471 | 2.0007 | 0.41201 | 1.9954
| Rel.Std.Dev. |  | 6.462% | 1.754% | 4.384% | 11.11% | 2.182% | 3.042% | 2.574% | 2.468% | 1.96% | 1.378% | 3.412% | 11.28% | 3.772% | 9.832% | 0.9801% | 5.003%
| PDM3 |  | -2.17% | -2.834% | -6.691% | -14% | 9.045% | -26.86% | -13.56% | 5.942% | 17.63% | -3.677% | 11.44% | -0.2635% | 9.09% | -45.88% | 11.81% | 6.075%
| Z-Score (Absolute) |  | 0.06 | 0.11 | 0.36 | 0.84 | 0.67 | 1.68 | 0.81 | 0.47 | 1.24 | 0.16 | 0.83 | 0.06 | 0.68 | 2.93 | 0.85 | 0.48
| Analytical results for Lu in OREAS 290 (Indicative Value 0.11 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.09 | 2.11 | 2.14 | 2.17 | [15 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [15 empty] | 2
|  |  | 0.25g | 0.25g | 0.5g | 0.5g | [15 empty] | 2
| 1 | 1 | 0.1 | 0.1 | 0.119 | 0.1 | [15 empty] | 1
| 1 | 2 | 0.11 | 0.1 | 0.108 | 0.1 | [15 empty] | 18
| 1 | 3 | 0.11 | 0.1 | 0.111 | 0.1 | [15 empty] | 16
| 1 | 4 | 0.11 | 0.09 | 0.115 | 0.1 | [15 empty] | 0.10592
| 1 | 5 | 0.12 | 0.11 | 0.116 | 0.1 | [15 empty] | 24
| 1 | 6 | 0.11 | 0.1 | 0.113 | 0.1
| Mean |  | 0.11 | 0.1 | 0.11367 | 0.1
| Median |  | 0.11 | 0.1 | 0.114 | 0.1
| Std Dev. |  | 0.0063246 | 0.0063246 | 0.0038816 | 1.5202e-17
| Rel.Std.Dev. |  | 5.75% | 6.325% | 3.415% | 1.52e-14%
| PDM3 |  | 3.855% | -5.586% | 7.317% | -5.586%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 1.17 | 0.67
| Analytical results for Mg in OREAS 290 (Certified Value 1.14 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 2
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 1.0605 | 1.0534 | 1.11 | 1.16 | 1.16 | 1.13 | 1.12 | 1.2 | 1.13 | 1.13 | 1.17 | 1.2469 | 1.157 | 1.28 | 1.02 | 1.08 | 0.623 | 1.235 | 1.16 | 1
| 1 | 2 | 1.0815 | 1.0538 | 1.12 | 1.12 | 1.17 | 1.14 | 1.12 | 1.2 | 1.16 | 1.14 | 1.18 | 1.2858 | 1.113 | 1.26 | 1.28 | 1.11 | 0.582 | 1.23 | 1.24 | 7
| 1 | 3 | 1.155 | 1.0411 | 1.13 | 1.15 | 1.16 | 1.16 | 1.14 | 1.21 | 1.15 | 1.11 | 1.19 | 1.2484 | 1.131 | 1.25 | 1.12 | 1.15 | 0.794 | 1.21 | 1.22 | 16
| 1 | 4 | 1.0815 | 1.0323 | 1.13 | 1.16 | 1.18 | 1.15 | 1.2 | 1.23 | 1.16 | 1.12 | 1.2 | 1.2424 | 1.164 | 1.27 | 1.39 | 1.15 | 0.815 | 1.23 | 1.33 | 1.145
| 1 | 5 | 1.092 | 1.061 | 1.13 | 1.17 | 1.18 | 1.13 | 1.14 | 1.24 | 1.14 | 1.14 | 1.22 | 1.2888 | 1.158 | 1.28 | 1.06 | 1.11 | 0.779 | 1.225 | 1.43 | 96
| 1 | 6 | 1.071 | 1.0691 | 1.14 | 1.17 | 1.15 | 1.14 | 1.13 | 1.23 | 1.14 | 1.14 | 1.21 | 1.2488 | 1.138 | 1.25 | 1.17 | 1.18 | 0.698 | 1.22 | 1.33
| Mean |  | 1.0903 | 1.0518 | 1.1267 | 1.155 | 1.1667 | 1.1417 | 1.1417 | 1.2183 | 1.1467 | 1.13 | 1.195 | 1.2602 | 1.1435 | 1.265 | 1.1733 | 1.13 | 0.71517 | 1.225 | 1.285
| Median |  | 1.0815 | 1.0536 | 1.13 | 1.16 | 1.165 | 1.14 | 1.135 | 1.22 | 1.145 | 1.135 | 1.195 | 1.2486 | 1.1475 | 1.265 | 1.145 | 1.13 | 0.7385 | 1.2275 | 1.285
| Std Dev. |  | 0.033479 | 0.013302 | 0.010328 | 0.018708 | 0.012111 | 0.01169 | 0.029944 | 0.017224 | 0.012111 | 0.012649 | 0.018708 | 0.02116 | 0.019644 | 0.013784 | 0.13967 | 0.036332 | 0.096725 | 0.0089443 | 0.096902
| Rel.Std.Dev. |  | 3.071% | 1.265% | 0.9167% | 1.62% | 1.038% | 1.024% | 2.623% | 1.414% | 1.056% | 1.119% | 1.566% | 1.679% | 1.718% | 1.09% | 11.9% | 3.215% | 13.52% | 0.7301% | 7.541%
| PDM3 |  | -4.779% | -8.141% | -1.599% | 0.8759% | 1.895% | -0.2886% | -0.2886% | 6.407% | 0.1481% | -1.308% | 4.369% | 10.06% | -0.1285% | 10.48% | 2.477% | -1.308% | -37.54% | 6.99% | 12.23%
| Z-Score (Absolute) |  | 1.43 | 2.4 | 0.51 | 0.21 | 0.51 | 0.13 | 0.13 | 1.81 | 0 | 0.42 | 1.22 | 2.87 | 0.08 | 2.99 | 0.67 | 0.42 | 10.91 | 1.98 | 3.5
| Analytical results for Mn in OREAS 290 (Certified Value 0.047 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.047006 | 0.042067 | 0.0446 | 0.0467 | 0.0455 | 0.0457 | 0.0479 | 0.0486 | 0.0398 | 0.0487 | 0.0468 | 0.050367 | 0.04973 | 0.0472 | 0.041 | 0.0498 | 0.036 | 0.045929 | 0.0486 | 1
| 1 | 2 | 0.04652 | 0.042143 | 0.0456 | 0.0448 | 0.046 | 0.0461 | 0.0471 | 0.049 | 0.0408 | 0.049 | 0.0469 | 0.049925 | 0.04876 | 0.0465 | 0.0512 | 0.0499 | 0.035 | 0.045853 | 0.0514 | 8
| 1 | 3 | 0.048955 | 0.042057 | 0.046 | 0.0457 | 0.0459 | 0.0457 | 0.0483 | 0.0492 | 0.0402 | 0.0485 | 0.0474 | 0.050365 | 0.04919 | 0.0462 | 0.0448 | 0.0529 | 0.04 | 0.045893 | 0.0522 | 16
| 1 | 4 | 0.046795 | 0.042044 | 0.046 | 0.0462 | 0.0466 | 0.0458 | 0.0488 | 0.0499 | 0.0406 | 0.0485 | 0.048 | 0.050983 | 0.04916 | 0.0467 | 0.055 | 0.0498 | 0.04 | 0.045615 | 0.056 | 0.047492
| 1 | 5 | 0.046351 | 0.043117 | 0.0458 | 0.0465 | 0.0462 | 0.0449 | 0.048 | 0.0502 | 0.0401 | 0.048 | 0.0483 | 0.050793 | 0.04999 | 0.0476 | 0.0424 | 0.0509 | 0.039 | 0.045838 | 0.0534 | 97
| 1 | 6 | 0.047081 | 0.043446 | 0.0455 | 0.0463 | 0.0456 | 0.0458 | 0.0475 | 0.0501 | 0.0398 | 0.0486 | 0.0486 | 0.050262 | 0.05009 | 0.0464 | 0.0465 | 0.0514 | 0.037 | 0.045764 | 0.0564
| Mean |  | 0.047118 | 0.042479 | 0.045583 | 0.046033 | 0.045967 | 0.045667 | 0.047933 | 0.0495 | 0.040217 | 0.04855 | 0.047667 | 0.050449 | 0.049487 | 0.046767 | 0.046817 | 0.050783 | 0.037833 | 0.045815 | 0.053
| Median |  | 0.046901 | 0.042105 | 0.0457 | 0.04625 | 0.04595 | 0.04575 | 0.04795 | 0.04955 | 0.04015 | 0.04855 | 0.0477 | 0.050366 | 0.04946 | 0.0466 | 0.04565 | 0.0504 | 0.038 | 0.045845 | 0.0528
| Std Dev. |  | 0.00094215 | 0.0006312 | 0.00052313 | 0.00069186 | 0.00040332 | 0.00040332 | 0.00059554 | 0.00065727 | 0.00041191 | 0.00032711 | 0.00074744 | 0.00038131 | 0.00052902 | 0.00053166 | 0.0053608 | 0.0012319 | 0.002137 | 0.00011261 | 0.0029421
| Rel.Std.Dev. |  | 2% | 1.486% | 1.148% | 1.503% | 0.8774% | 0.8832% | 1.242% | 1.328% | 1.024% | 0.6738% | 1.568% | 0.7558% | 1.069% | 1.137% | 11.45% | 2.426% | 5.648% | 0.2458% | 5.551%
| PDM3 |  | -0.7875% | -10.56% | -4.019% | -3.071% | -3.212% | -3.843% | 0.9295% | 4.228% | -15.32% | 2.228% | 0.368% | 6.227% | 4.2% | -1.527% | -1.422% | 6.931% | -20.34% | -3.53% | 11.6%
| Z-Score (Absolute) |  | 0.16 | 2.37 | 0.67 | 0.43 | 0.46 | 0.63 | 0.61 | 1.47 | 3.61 | 0.95 | 0.46 | 1.99 | 1.46 | 0.03 | 0 | 2.17 | 4.91 | 0.55 | 3.38
| Analytical results for Mo in OREAS 290 (Certified Value 1.09 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 2
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 1.221 | 0.27 | 1.12 | 0.96 | 1.01 | 0.97 | 1.42 | 1.02 | < 1 | 1.21 | 1 | 1.1103 | 1.16 | 1.3 | 1 | 0.9 | < 1 | < 2 | 1.3 | 1
| 1 | 2 | 1.023 | 0.38 | 1.03 | 0.95 | 1 | 1.04 | 1.34 | 1.28 | 1 | 1.24 | 1 | 1.1148 | 1.12 | 1.2 | 1 | 0.96 | < 1 | < 2 | 1.3 | 9
| 1 | 3 | 1.012 | 0.68 | 1.22 | 1.26 | 1.03 | 1.12 | 1.18 | 1.15 | < 1 | 1.25 | < 1 | 1.0014 | 1.1 | 1.2 | < 1 | 1.01 | < 1 | < 2 | 1.1 | 16
| 1 | 4 | < 1 | 0.38 | 1.16 | 0.97 | 1.06 | 1.02 | 1.03 | 1.04 | < 1 | 1.23 | < 1 | 1.0158 | 1.17 | 1.1 | 1 | 1.05 | 1.1 | < 2 | 1.1 | 1.0946
| 1 | 5 | < 1 | 1.35 | 1.08 | 1 | 1.04 | 0.97 | 1.05 | 1.02 | < 1 | 1.29 | < 1 | 1.0469 | 1.11 | 1.2 | < 1 | 0.97 | 1.2 | < 2 | 1.2 | 98
| 1 | 6 | < 1 | 0.71 | 1.09 | 1.16 | 1.03 | 1 | 1.12 | 1.04 | < 1 | 1.28 | < 1 | 0.97375 | 1.17 | 1.1 | < 1 | 0.92 | < 1 | < 2 | 1.1
| Mean |  | 1.0853 | 0.62833 | 1.1167 | 1.05 | 1.0283 | 1.02 | 1.19 | 1.0917 | 1 | 1.25 | 1 | 1.0438 | 1.1383 | 1.1833 | 1 | 0.96833 | 1.15 |  | 1.1833
| Median |  | 1.023 | 0.53 | 1.105 | 0.985 | 1.03 | 1.01 | 1.15 | 1.04 | 1 | 1.245 | 1 | 1.0314 | 1.14 | 1.2 | 1 | 0.965 | 1.15 |  | 1.15
| Std Dev. |  | 0.11762 | 0.39544 | 0.066533 | 0.129 | 0.02137 | 0.056214 | 0.15849 | 0.10439 |  | 0.030332 | 0 | 0.058257 | 0.031885 | 0.075277 | 0 | 0.055648 | 0.070711 |  | 0.098319
| Rel.Std.Dev. |  | 10.84% | 62.94% | 5.958% | 12.29% | 2.078% | 5.511% | 13.32% | 9.562% |  | 2.427% | 0% | 5.581% | 2.801% | 6.361% | 0% | 5.747% | 6.149% |  | 8.309%
| PDM3 |  | -0.8496% | -42.6% | 2.013% | -4.077% | -6.057% | -6.818% | 8.712% | -0.271% | -8.645% | 14.19% | -8.645% | -4.641% | 3.992% | 8.103% | -8.645% | -11.54% | 5.058% |  | 8.103%
| Z-Score (Absolute) |  | 1.79 | 2.96 | 0.52 | 0.04 | 0.11 | 0.17 | 1.04 | 0.34 | NA | 1.47 | NA | 0 | 0.67 | 1 | 2.1 | 0.54 | 2.33 | 0.31 | 1
| Analytical results for Na in OREAS 290 (Certified Value 0.047 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.041428 | 0.048539 | 0.044 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.060061 | 0.049 | 0.09 | 0.04 | 0.048 | 0.0283 | 0.0465 | 0.0468 | 1
| 1 | 2 | 0.042753 | 0.050984 | 0.045 | 0.04 | 0.05 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.061561 | 0.047 | 0.09 | 0.05 | 0.051 | 0.0276 | 0.047275 | 0.045 | #N/A
| 1 | 3 | 0.04641 | 0.049746 | 0.045 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.060237 | 0.049 | 0.09 | 0.04 | 0.054 | 0.033 | 0.0465 | 0.0388 | 16
| 1 | 4 | 0.04305 | 0.048913 | 0.045 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05994 | 0.049 | 0.09 | 0.05 | 0.052 | 0.0325 | 0.0465 | 0.0465 | 0.046558
| 1 | 5 | 0.04406 | 0.051681 | 0.045 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.061411 | 0.049 | 0.09 | 0.04 | 0.057 | 0.0312 | 0.0465 | 0.0487 | 99
| 1 | 6 | 0.043296 | 0.051189 | 0.046 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.06229 | 0.049 | 0.09 | 0.05 | 0.054 | 0.0307 | 0.0465 | 0.0461
| Mean |  | 0.043499 | 0.050175 | 0.045 | 0.046667 | 0.05 | 0.04 | 0.05 | 0.048333 | 0.04 | 0.041667 | 0.05 | 0.060917 | 0.048667 | 0.09 | 0.045 | 0.052667 | 0.03055 | 0.046629 | 0.045317
| Median |  | 0.043173 | 0.050365 | 0.045 | 0.05 | 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.060824 | 0.049 | 0.09 | 0.045 | 0.053 | 0.03095 | 0.0465 | 0.0463
| Std Dev. |  | 0.0016659 | 0.0012964 | 0.00063246 | 0.005164 | 7.6012e-18 | 0 | 7.6012e-18 | 0.0040825 | 0 | 0.0040825 | 7.6012e-18 | 0.00096891 | 0.0008165 | 1.5202e-17 | 0.0054772 | 0.0030768 | 0.0021916 | 0.00031639 | 0.0034126
| Rel.Std.Dev. |  | 3.83% | 2.584% | 1.405% | 11.07% | 1.52e-14% | 0% | 1.52e-14% | 8.447% | 0% | 9.798% | 1.52e-14% | 1.591% | 1.678% | 1.689e-14% | 12.17% | 5.842% | 7.174% | 0.6785% | 7.531%
| PDM3 |  | -6.569% | 7.77% | -3.346% | 0.2339% | 7.394% | -14.09% | 7.394% | 3.814% | -14.09% | -10.51% | 7.394% | 30.84% | 4.53% | 93.31% | -3.346% | 13.12% | -34.38% | 0.1534% | -2.666%
| Z-Score (Absolute) |  | 0.64 | 0.71 | 0.34 | 0 | 0.67 | 1.35 | 0.67 | 0.34 | 1.35 | 1.01 | 0.67 | 2.88 | 0.4 | 8.77 | 0.34 | 1.21 | 3.26 | 0.01 | 0.27
| Analytical results for Nb in OREAS 290 (Certified Value 0.54 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [7 empty] | 2
| 1 | 1 | 0.72 | 0.66 | 0.47 | 0.42 | 0.52 | 0.36 | 0.73142 | 0.87 | 0.4 | < 0.1 | < 5 | 0.43 | [7 empty] | 1
| 1 | 2 | 0.65 | 0.55 | 0.53 | 0.48 | 0.5 | 0.36 | 0.75806 | 0.79 | 0.4 | < 0.1 | < 5 | 0.38 | [7 empty] | 10
| 1 | 3 | 0.57 | 0.58 | 0.56 | 0.5 | 0.56 | 0.36 | 0.78731 | 0.84 | 0.4 | 0.1 | < 5 | 0.44 | [7 empty] | 16
| 1 | 4 | 0.6 | 0.61 | 0.59 | 0.55 | 0.54 | 0.36 | 0.74167 | 0.85 | 0.4 | < 0.1 | < 5 | 0.46 | [7 empty] | 0.53966
| 1 | 5 | 0.61 | 0.61 | 0.58 | 0.4 | 0.57 | 0.36 | 0.74492 | 0.86 | 0.4 | < 0.1 | < 5 | 0.39 | [7 empty] | 100
| 1 | 6 | 0.64 | 0.59 | 0.53 | 0.47 | 0.58 | 0.39 | 0.78008 | 0.89 | 0.4 | < 0.1 | < 5 | 0.36
| Mean |  | 0.63167 | 0.6 | 0.54333 | 0.47 | 0.545 | 0.365 | 0.75724 | 0.85 | 0.4 | 0.1 |  | 0.41
| Median |  | 0.625 | 0.6 | 0.545 | 0.475 | 0.55 | 0.36 | 0.75149 | 0.855 | 0.4 | 0.1 |  | 0.41
| Std Dev. |  | 0.051929 | 0.036878 | 0.043665 | 0.054406 | 0.030822 | 0.012247 | 0.022304 | 0.034059 | 6.0809e-17 |  |  | 0.038987
| Rel.Std.Dev. |  | 8.221% | 6.146% | 8.037% | 11.58% | 5.655% | 3.355% | 2.945% | 4.007% | 1.52e-14% |  |  | 9.509%
| PDM3 |  | 17.05% | 11.18% | 0.6816% | -12.91% | 0.9904% | -32.36% | 40.32% | 57.51% | -25.88% | -81.47% |  | -24.03%
| Z-Score (Absolute) |  | 0.3 | 0.14 | 0.15 | 0.52 | 0.14 | 1.06 | 0.94 | 1.41 | NA | NA | 9.81 | 0.83
Results from laboratories 2.15 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 290 (Indicative Value 19.7 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 0.5g | 0.5g | [17 empty] | 1
| 1 | 1 | 20.927 | 19.2 | [17 empty] | 1
| 1 | 2 | 19.306 | 19.8 | [17 empty] | 19
| 1 | 3 | 19.741 | 19.3 | [17 empty] | 16
| 1 | 4 | 20.379 | 18.9 | [17 empty] | 19.7263
| 1 | 5 | 20.798 | 18.9 | [17 empty] | 25
| 1 | 6 | 20.465 | 19
| Mean |  | 20.2693 | 19.1833
| Median |  | 20.422 | 19.1
| Std Dev. |  | 0.62761 | 0.34303
| Rel.Std.Dev. |  | 3.096% | 1.788%
| PDM3 |  | 2.753% | -2.753%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 290 (Certified Value 68 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 1
| 1 | 1 | 76.4348 | 79.49 | 67.1 | 67.1 | 70.4 | 66.8 | 73.6 | 67.6 | 56 | 64 | 65 | 70.7983 | 70.8 | 76 | 54 | 65.8 | 57 | 67.976 | 70.4 | 1
| 1 | 2 | 76.5555 | 79.97 | 66.1 | 65.4 | 69.4 | 70.8 | 73.1 | 68.2 | 58 | 64.3 | 64 | 70.5423 | 72 | 75 | 67 | 63.7 | 54 | 67.9654 | 70.9 | 11
| 1 | 3 | 77.4008 | 81.64 | 66.6 | 71.4 | 70.3 | 73.9 | 68.6 | 69.3 | 56 | 64.8 | 65 | 69.2783 | 71.4 | 75 | 58 | 66.2 | 62 | 67.8533 | 63.5 | 16
| 1 | 4 | 75.7103 | 80.63 | 67.8 | 66.8 | 70.8 | 72.2 | 64.6 | 69.2 | 57 | 64.6 | 65 | 70.0122 | 72.2 | 76 | 72 | 66.7 | 63 | 67.5465 | 70 | 67.7149
| 1 | 5 | 77.28 | 80.43 | 67.8 | 69.2 | 71.5 | 71.7 | 64.8 | 68.9 | 57 | 63.9 | 67 | 69.6773 | 73.5 | 76 | 55 | 66.4 | 62 | 67.8396 | 68.4 | 101
| 1 | 6 | 77.28 | 82.35 | 67.9 | 76.6 | 67.9 | 72.2 | 69.4 | 68.4 | 56 | 64.8 | 66 | 69.9 | 72.2 | 75 | 61 | 65.1 | 57 | 67.395 | 64.4
| Mean |  | 76.7769 | 80.7517 | 67.2167 | 69.4167 | 70.05 | 71.2667 | 69.0167 | 68.6 | 56.6667 | 64.4 | 65.3333 | 70.0347 | 72.0167 | 75.5 | 61.1667 | 65.65 | 59.1667 | 67.7626 | 67.9333
| Median |  | 76.9177 | 80.53 | 67.45 | 68.15 | 70.35 | 71.95 | 69 | 68.65 | 56.5 | 64.45 | 65 | 69.9561 | 72.1 | 75.5 | 59.5 | 66 | 59.5 | 67.8464 | 69.2
| Std Dev. |  | 0.66321 | 1.0644 | 0.74677 | 4.097 | 1.2566 | 2.4097 | 3.8804 | 0.65422 | 0.8165 | 0.39497 | 1.0328 | 0.55845 | 0.90866 | 0.54772 | 7.0828 | 1.1041 | 3.656 | 0.23776 | 3.2098
| Rel.Std.Dev. |  | 0.8638% | 1.318% | 1.111% | 5.902% | 1.794% | 3.381% | 5.622% | 0.9537% | 1.441% | 0.6133% | 1.581% | 0.7974% | 1.262% | 0.7255% | 11.58% | 1.682% | 6.179% | 0.3509% | 4.725%
| PDM3 |  | 13.38% | 19.25% | -0.7357% | 2.513% | 3.448% | 5.245% | 1.922% | 1.307% | -16.32% | -4.895% | -3.517% | 3.426% | 6.353% | 11.5% | -9.67% | -3.049% | -12.62% | 0.07053% | 0.3226%
| Z-Score (Absolute) |  | 1.87 | 2.78 | 0.32 | 0.19 | 0.33 | 0.61 | 0.1 | 0 | 2.73 | 0.96 | 0.75 | 0.33 | 0.78 | 1.58 | 1.7 | 0.67 | 2.16 | 0.19 | 0.15
| Analytical results for P in OREAS 290 (Certified Value 0.05 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.04824 | 0.05036 | 0.048 | 0.052 | 0.049 | 0.05 | 0.051 | 0.0531 | 0.05 | 0.0503 | 0.05 | 0.051131 | 0.0488 | 0.0558 | 0.0412 | 0.034 | 0.0469 | 0.050873 | 0.0546 | 1
| 1 | 2 | 0.048253 | 0.051662 | 0.048 | 0.05 | 0.05 | 0.05 | 0.051 | 0.052 | 0.05 | 0.0491 | 0.05 | 0.051027 | 0.0477 | 0.0553 | 0.0515 | 0.034 | 0.043 | 0.05059 | 0.053 | 12
| 1 | 3 | 0.047323 | 0.049449 | 0.049 | 0.051 | 0.05 | 0.05 | 0.051 | 0.0531 | 0.05 | 0.0489 | 0.05 | 0.051923 | 0.0486 | 0.054 | 0.0446 | 0.039 | 0.052 | 0.050503 | 0.0621 | 16
| 1 | 4 | 0.04757 | 0.052272 | 0.049 | 0.051 | 0.051 | 0.05 | 0.052 | 0.0534 | 0.05 | 0.0502 | 0.05 | 0.052382 | 0.0504 | 0.0556 | 0.0553 | 0.035 | 0.0531 | 0.050192 | 0.0714 | 0.050105
| 1 | 5 | 0.048559 | 0.050079 | 0.049 | 0.052 | 0.051 | 0.049 | 0.052 | 0.0542 | 0.05 | 0.0492 | 0.06 | 0.051968 | 0.0511 | 0.0564 | 0.0426 | 0.033 | 0.0526 | 0.050411 | 0.062 | 102
| 1 | 6 | 0.048295 | 0.050508 | 0.049 | 0.052 | 0.049 | 0.05 | 0.051 | 0.0527 | 0.06 | 0.0497 | 0.05 | 0.052001 | 0.0502 | 0.0546 | 0.0464 | 0.031 | 0.0486 | 0.050175 | 0.0676
| Mean |  | 0.04804 | 0.050722 | 0.048667 | 0.051333 | 0.05 | 0.049833 | 0.051333 | 0.053083 | 0.051667 | 0.049567 | 0.051667 | 0.051739 | 0.049467 | 0.055283 | 0.046933 | 0.034333 | 0.049367 | 0.050457 | 0.061783
| Median |  | 0.048246 | 0.050434 | 0.049 | 0.0515 | 0.05 | 0.05 | 0.051 | 0.0531 | 0.05 | 0.04945 | 0.05 | 0.051946 | 0.0495 | 0.05545 | 0.0455 | 0.034 | 0.0503 | 0.050457 | 0.06205
| Std Dev. |  | 0.00048059 | 0.0010485 | 0.0005164 | 0.0008165 | 0.00089443 | 0.00040825 | 0.0005164 | 0.00073052 | 0.0040825 | 0.00059217 | 0.0040825 | 0.00053741 | 0.0012956 | 0.00086352 | 0.0054449 | 0.0026583 | 0.0039632 | 0.0002627 | 0.0071443
| Rel.Std.Dev. |  | 1% | 2.067% | 1.061% | 1.591% | 1.789% | 0.8192% | 1.006% | 1.376% | 7.902% | 1.195% | 7.902% | 1.039% | 2.619% | 1.562% | 11.6% | 7.743% | 8.028% | 0.5206% | 11.56%
| PDM3 |  | -4.123% | 1.23% | -2.872% | 2.451% | -0.2105% | -0.5431% | 2.451% | 5.943% | 3.116% | -1.075% | 3.116% | 3.26% | -1.275% | 10.33% | -6.331% | -31.48% | -1.474% | 0.702% | 23.31%
| Z-Score (Absolute) |  | 1.35 | 0.15 | 1 | 0.49 | 0.25 | 0.35 | 0.49 | 1.46 | 0.67 | 0.5 | 0.67 | 0.71 | 0.55 | 2.69 | 1.96 | 8.99 | 0.61 | 0 | 6.31
| Analytical results for Pb in OREAS 290 (Certified Value 13.6 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 2
| 1 | 1 | 14.41 | 1.16 | 11.89 | 13.1 | 13.3 | 13.5 | 13 | 13.9 | 11 | 12.9 | 13 | 14.3176 | 14.3 | 12.8 | 11 | 9.9 | 14 | 14.115 | 14.7 | 1
| 1 | 2 | 14.52 | 0.42 | 12.82 | 12.6 | 13.7 | 14 | 12.6 | 13.8 | 12 | 12.6 | 13 | 14.5245 | 14.4 | 13 | 15 | 9.4 | 15 | 14.26 | 16.5 | 13
| 1 | 3 | 13.64 | 2.11 | 11.27 | 14.8 | 13.7 | 14.4 | 12.5 | 14.1 | 12 | 12.2 | 13 | 13.6898 | 14.5 | 13 | 13 | 12.2 | 14 | 13.1608 | 14.9 | 16
| 1 | 4 | 14.08 | 3.23 | 12.39 | 13 | 13.9 | 14 | 11.8 | 14.3 | 12 | 11.9 | 14 | 14.0154 | 14.4 | 12.4 | 15 | 10.3 | 14 | 13.005 | 14.8 | 13.5606
| 1 | 5 | 14.41 | 1.43 | 12.44 | 13.6 | 13.9 | 13.9 | 10.4 | 14 | 13 | 12.8 | 13 | 14.0117 | 14.3 | 12.8 | 11 | 9.6 | 20 | 13.215 | 15.1 | 103
| 1 | 6 | 13.86 | 1 | 12 | 15.3 | 13.1 | 14 | 11.9 | 13.8 | 13 | 12.9 | 15 | 14.7372 | 14.4 | 12.8 | 12 | 9.1 | 14 | 13.4085 | 14.9
| Mean |  | 14.1533 | 1.5583 | 12.135 | 13.7333 | 13.6 | 13.9667 | 12.0333 | 13.9833 | 12.1667 | 12.55 | 13.5 | 14.216 | 14.3833 | 12.8 | 12.8333 | 10.0833 | 15.1667 | 13.5274 | 15.15
| Median |  | 14.245 | 1.295 | 12.195 | 13.35 | 13.7 | 14 | 12.2 | 13.95 | 12 | 12.7 | 13 | 14.1665 | 14.4 | 12.8 | 12.5 | 9.75 | 14 | 13.3117 | 14.9
| Std Dev. |  | 0.35246 | 0.98808 | 0.53921 | 1.0801 | 0.32863 | 0.28752 | 0.91797 | 0.19408 | 0.75277 | 0.41352 | 0.83666 | 0.38367 | 0.075277 | 0.21909 | 1.8348 | 1.1161 | 2.4014 | 0.52934 | 0.67454
| Rel.Std.Dev. |  | 2.49% | 63.41% | 4.443% | 7.865% | 2.416% | 2.059% | 7.629% | 1.388% | 6.187% | 3.295% | 6.197% | 2.699% | 0.5234% | 1.712% | 14.3% | 11.07% | 15.83% | 3.913% | 4.452%
| PDM3 |  | 4.371% | -88.51% | -10.51% | 1.273% | 0.2902% | 2.994% | -11.26% | 3.117% | -10.28% | -7.453% | -0.4472% | 4.833% | 6.067% | -5.609% | -5.363% | -25.64% | 11.84% | -0.2453% | 11.72%
| Z-Score (Absolute) |  | 0.48 | 10.37 | 1.26 | 0.11 | 0 | 0.32 | 1.35 | 0.33 | NA | 0.9 | NA | 0.53 | 0.67 | 0.69 | NA | 3.03 | NA | 0.06 | 1.33
Results from laboratories 2.10, 2.12, 2.16 and 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 290 (Indicative Value < 10 ppb) | [20 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.13 | 2.14 | [17 empty] | ppb
|  |  | AR*OES/MS | AR*MS | [17 empty] | 1
|  |  | 01g | 0.5g | [17 empty] | 1
| 1 | 1 | < 10 | 1 | [17 empty] | 1
| 1 | 2 | < 10 | < 1 | [17 empty] | 1
| 1 | 3 | < 10 | < 1 | [17 empty] | 16
| 1 | 4 | < 10 | 1 | [17 empty] | < 10
| 1 | 5 | < 10 | < 1 | [17 empty] | 26
| 1 | 6 | < 10 | 1
| Mean |  |  | 1
| Median |  |  | 1
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 290 (Indicative Value 5.3 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 5.643 | 5 | [17 empty] | 1
| 1 | 2 | 5.612 | 5 | [17 empty] | 21
| 1 | 3 | 5.622 | 5 | [17 empty] | 16
| 1 | 4 | 5.658 | 5 | [17 empty] | 5.3024
| 1 | 5 | 5.604 | 5 | [17 empty] | 27
| 1 | 6 | 5.59 | 4.9
| Mean |  | 5.6215 | 4.9833
| Median |  | 5.617 | 5
| Std Dev. |  | 0.025233 | 0.040825
| Rel.Std.Dev. |  | 0.4489% | 0.8192%
| PDM3 |  | 6.018% | -6.018%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pt in OREAS 290 (Indicative Value 2 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 2.13 | 2.14 | 2.20 | [16 empty] | ppb
|  |  | AR*OES/MS | AR*MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.5g | 0.1g | [16 empty] | 2
| 1 | 1 | < 5 | 3 | 1 | [16 empty] | 1
| 1 | 2 | < 5 | 1 | 3 | [16 empty] | 1
| 1 | 3 | < 5 | 2 | 2 | [16 empty] | 16
| 1 | 4 | < 5 | 3 | 2 | [16 empty] | 2
| 1 | 5 | < 5 | 2 | 2 | [16 empty] | 23
| 1 | 6 | < 5 | 2 | 1
| Mean |  |  | 2.1667 | 1.8333
| Median |  |  | 2 | 2
| Std Dev. |  |  | 0.75277 | 0.75277
| Rel.Std.Dev. |  |  | 34.74% | 41.06%
| PDM3 |  |  | 8.333% | -8.333%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.67
| Analytical results for Rb in OREAS 290 (Certified Value 81 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.20 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.1g | [8 empty] | 1
| 1 | 1 | 79.5 | 81.2 | 75.3 | 91.7 | 80.6 | 74.7 | 83.7866 | 83.63 | 85.6 | 71 | 84.3 | [8 empty] | 1
| 1 | 2 | 78.3 | 80.9 | 81.4 | 80.1 | 81.6 | 76.5 | 85.5284 | 83.07 | 88.1 | 71.5 | 80.4 | [8 empty] | 44
| 1 | 3 | 78.3 | 83.5 | 84 | 83.5 | 82.2 | 73.7 | 83.6279 | 84.28 | 87.6 | 72 | 77.5 | [8 empty] | 16
| 1 | 4 | 73.2 | 83.2 | 79.8 | 47.1 | 82.6 | 73 | 82.8392 | 84.23 | 88.5 | 71.1 | 86.3 | [8 empty] | 80.8171
| 1 | 5 | 77.1 | 83.1 | 79.5 | 63.9 | 83.9 | 77 | 84.5729 | 84.6 | 87.6 | 73.1 | 79.9 | [8 empty] | 104
| 1 | 6 | 87.6 | 81.5 | 79.7 | 80.6 | 80.9 | 77.1 | 85.3339 | 84.5 | 85.3 | 73.2 | 76.5
| Mean |  | 79 | 82.2333 | 79.95 | 74.4833 | 81.9667 | 75.3333 | 84.2815 | 84.0517 | 87.1167 | 71.9833 | 80.8167
| Median |  | 78.3 | 82.3 | 79.75 | 80.35 | 81.9 | 75.6 | 84.1798 | 84.255 | 87.6 | 71.75 | 80.15
| Std Dev. |  | 4.7422 | 1.1553 | 2.8431 | 16.1773 | 1.2111 | 1.7761 | 1.0488 | 0.58772 | 1.3378 | 0.9704 | 3.817
| Rel.Std.Dev. |  | 6.003% | 1.405% | 3.556% | 21.72% | 1.478% | 2.358% | 1.244% | 0.6992% | 1.536% | 1.348% | 4.723%
| PDM3 |  | -2.248% | 1.752% | -1.073% | -7.837% | 1.422% | -6.785% | 4.287% | 4.002% | 7.795% | -10.93% | -0.0005435%
| Z-Score (Absolute) |  | 0.38 | 0.3 | 0.18 | 1.32 | 0.24 | 1.14 | 0.72 | 0.67 | 1.31 | 1.84 | 0
| Analytical results for Re in OREAS 290 (Certified Value < 0.001 ppm) | [20 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.1 | < 0.002 | < 0.05 | < 0.001 | < 0.05 | < 0.001 | [9 empty] | 1
| 1 | 2 | < 0.001 | 0.001 | 0.001 | 0.001 | < 0.1 | < 0.002 | < 0.05 | < 0.001 | < 0.05 | < 0.001 | [9 empty] | 45
| 1 | 3 | 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.1 | < 0.002 | < 0.05 | < 0.001 | < 0.05 | < 0.001 | [9 empty] | 16
| 1 | 4 | < 0.001 | < 0.001 | < 0.001 | 0.001 | < 0.1 | < 0.002 | < 0.05 | < 0.001 | < 0.05 | < 0.001 | [9 empty] | < 0.001
| 1 | 5 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.1 | < 0.002 | < 0.05 | < 0.001 | < 0.05 | < 0.001 | [9 empty] | 105
| 1 | 6 | < 0.001 | 0.001 | 0.001 | < 0.001 | < 0.1 | < 0.002 | < 0.05 | < 0.001 | < 0.05 | < 0.001
| Mean |  | 0.001 | 0.001 | 0.001 | 0.001
| Median |  | 0.001 | 0.001 | 0.001 | 0.001
| Std Dev. |  |  | 0 | 0 | 0
| Rel.Std.Dev. |  |  | 0% | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 0.58 | 0.19 | 0.39 | 0.19 | 113.67 | 0.39 | 55.87 | 0.77 | 55.87 | 0.77
| Analytical results for S in OREAS 290 (Certified Value 0.533 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.55776 | 0.55419 | 0.51 | 0.54 | 0.53 | 0.55 | 0.54 | 0.51 | 0.51 | 0.55 | 0.52 | 0.51882 | 0.51 | 0.53 | 0.3948 | 0.545 | 0.534 | 0.512 | 0.642 | 1
| 1 | 2 | 0.56336 | 0.56932 | 0.51 | 0.52 | 0.53 | 0.55 | 0.53 | 0.5 | 0.52 | 0.54 | 0.53 | 0.53183 | 0.52 | 0.52 | 0.4936 | 0.545 | 0.525 | 0.521 | 0.693 | 30
| 1 | 3 | 0.5488 | 0.54892 | 0.51 | 0.53 | 0.53 | 0.55 | 0.54 | 0.5 | 0.51 | 0.53 | 0.53 | 0.51546 | 0.53 | 0.52 | 0.4283 | 0.539 | 0.551 | 0.4985 | 0.65 | 16
| 1 | 4 | 0.55104 | 0.57426 | 0.52 | 0.54 | 0.54 | 0.55 | 0.56 | 0.51 | 0.53 | 0.54 | 0.53 | 0.52827 | 0.52 | 0.53 | 0.5341 | 0.543 | 0.556 | 0.503 | 0.732 | 0.53261
| 1 | 5 | 0.57008 | 0.55069 | 0.52 | 0.54 | 0.54 | 0.54 | 0.54 | 0.52 | 0.53 | 0.54 | 0.55 | 0.51321 | 0.53 | 0.53 | 0.4087 | 0.545 | 0.557 | 0.512 | 0.789 | 106
| 1 | 6 | 0.56896 | 0.55348 | 0.51 | 0.54 | 0.53 | 0.55 | 0.53 | 0.51 | 0.52 | 0.55 | 0.53 | 0.5199 | 0.53 | 0.52 | 0.4443 | 0.541 | 0.537 | 0.503 | 0.804
| Mean |  | 0.56 | 0.55848 | 0.51333 | 0.535 | 0.53333 | 0.54833 | 0.54 | 0.50833 | 0.52 | 0.54167 | 0.53167 | 0.52125 | 0.52333 | 0.525 | 0.45063 | 0.543 | 0.54333 | 0.50825 | 0.71833
| Median |  | 0.56056 | 0.55384 | 0.51 | 0.54 | 0.53 | 0.55 | 0.54 | 0.51 | 0.52 | 0.54 | 0.53 | 0.51936 | 0.525 | 0.525 | 0.4363 | 0.544 | 0.544 | 0.5075 | 0.7125
| Std Dev. |  | 0.008988 | 0.010603 | 0.005164 | 0.0083666 | 0.005164 | 0.0040825 | 0.010954 | 0.0075277 | 0.0089443 | 0.0075277 | 0.0098319 | 0.007307 | 0.008165 | 0.0054772 | 0.053343 | 0.0025298 | 0.013186 | 0.0082568 | 0.068786
| Rel.Std.Dev. |  | 1.605% | 1.899% | 1.006% | 1.564% | 0.9682% | 0.7445% | 2.029% | 1.481% | 1.72% | 1.39% | 1.849% | 1.402% | 1.56% | 1.043% | 11.84% | 0.4659% | 2.427% | 1.625% | 9.576%
| PDM3 |  | 5.143% | 4.857% | -3.619% | 0.4494% | 0.1365% | 2.953% | 1.388% | -4.557% | -2.367% | 1.701% | -0.1764% | -2.132% | -1.741% | -1.428% | -15.39% | 1.951% | 2.014% | -4.573% | 34.87%
| Z-Score (Absolute) |  | 1.49 | 1.4 | 1.12 | 0.09 | 0 | 0.84 | 0.37 | 1.39 | 0.74 | 0.46 | 0.09 | 0.67 | 0.56 | 0.46 | 4.61 | 0.54 | 0.56 | 1.4 | 10.32
| Analytical results for Sb in OREAS 290 (Certified Value 0.745 wt.%) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*MS | AR*OES | AR*AAS | HBr*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*AAS | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.2g | 0.5g |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 01g | 01g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.1g | 3
| 1 | 1 | 0.82382 | 0.74531 | 0.91 | 0.74 | 0.733 | 0.671 | 0.676 | 0.643 | 0.78132 | 0.568 | 0.74447 | 0.73279 | > 0.8204 | 0.489 | > 0.5 | 0.828 | 0.84 | 0.68023 | 0.814 | 1
| 1 | 2 | 0.78587 | 0.7502 | 0.9 | 0.71 | 0.701 | 0.679 | 0.671 | 0.633 | 0.79933 | 0.59 | 0.7563 | 0.72386 | > 0.8204 | 0.49 | > 0.5 | 0.831 | 0.81 | 0.67747 | 0.883 | 26
| 1 | 3 | 0.81962 | 0.74341 | 0.89 | 0.82 | 0.717 | 0.681 | 0.682 | 0.649 | 0.81574 | 0.574 | 0.73979 | 0.71524 | > 0.8188 | 0.471 | > 0.5 | 0.829 | 0.84 | 0.67143 | 0.795 | 16
| 1 | 4 | 0.80186 | 0.76612 | 0.89 | 0.71 | 0.721 | 0.687 | 0.682 | 0.775 | 0.82103 | 0.578 | 0.73532 | 0.73441 | > 0.8238 | 0.466 | > 0.5 | 0.829 | 0.85 | 0.67622 | 0.808 | 0.74463
| 1 | 5 | 0.79862 | 0.74995 | 0.89 | 0.82 | 0.729 | 0.68 | 0.663 | 0.643 | 0.82093 | 0.578 | 0.74554 | 0.72183 | > 0.8333 | 0.492 | > 0.5 | 0.836 | 0.83 | 0.6838 | 0.809 | 107
| 1 | 6 | 0.80033 | 0.75023 | 0.91 | 0.69 | 0.731 | 0.673 | 0.675 | 0.629 | 0.79548 | 0.574 | 0.75137 | 0.72256 | > 0.8121 | 0.466 | > 0.5 | 0.805 | 0.81 | 0.68237 | 0.817
| Mean |  | 0.80502 | 0.75087 | 0.89833 | 0.74833 | 0.722 | 0.6785 | 0.67483 | 0.662 | 0.80564 | 0.577 | 0.74547 | 0.72512 |  | 0.479 |  | 0.82633 | 0.83 | 0.67859 | 0.821
| Median |  | 0.80109 | 0.75007 | 0.895 | 0.725 | 0.725 | 0.6795 | 0.6755 | 0.643 | 0.80753 | 0.576 | 0.74501 | 0.72321 |  | 0.48 |  | 0.829 | 0.835 | 0.67885 | 0.8115
| Std Dev. |  | 0.014187 | 0.0080105 | 0.0098319 | 0.057764 | 0.011983 | 0.0057879 | 0.0071949 | 0.055839 | 0.016167 | 0.0073485 | 0.0075934 | 0.0072336 |  | 0.012586 |  | 0.010838 | 0.016733 | 0.0045226 | 0.031299
| Rel.Std.Dev. |  | 1.762% | 1.067% | 1.094% | 7.719% | 1.66% | 0.853% | 1.066% | 8.435% | 2.007% | 1.274% | 1.019% | 0.9976% |  | 2.627% |  | 1.312% | 2.016% | 0.6665% | 3.812%
| PDM3 |  | 8.11% | 0.8383% | 20.64% | 0.4978% | -3.039% | -8.881% | -9.373% | -11.1% | 8.194% | -22.51% | 0.1127% | -2.62% |  | -35.67% |  | 10.97% | 11.47% | -8.869% | 10.26%
| Z-Score (Absolute) |  | 0.6 | 0.05 | 1.54 | 0.03 | 0.24 | 0.67 | 0.71 | 0.84 | 0.61 | 1.7 | 0 | 0.2 | NA | 2.68 | NA | 0.81 | 0.85 | 0.67 | 0.76
| Analytical results for Sc in OREAS 290 (Certified Value 6.78 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.1g |  |  |  | 3
| 1 | 1 | 6.528 | 6.6 | 7.3 | 7.1 | 7 | 7 | 7.2 | 5.8 | 7.1 | 6.7 | 6.2334 | 7.47 | 7 | 5.8 | 5.6 | 7 |  |  |  | 1
| 1 | 2 | 6.66 | 6.6 | 7.2 | 7 | 7.7 | 6.8 | 7.5 | 5.9 | 7 | 6.6 | 6.4099 | 7.71 | 7 | 7.3 | 5.3 | 7 |  |  |  | 47
| 1 | 3 | 7.116 | 6.7 | 6.8 | 7.2 | 7.9 | 6.9 | 7.7 | 5.8 | 6.8 | 6.7 | 5.9944 | 7.52 | 7 | 6.4 | 6.3 | 7 |  |  |  | 16
| 1 | 4 | 6.636 | 6.7 | 6.4 | 7.1 | 7.5 | 6.7 | 7.7 | 5.9 | 6.9 | 6.8 | 6.0422 | 7.64 | 7 | 7.9 | 5.9 | 7 |  |  |  | 6.7834
| 1 | 5 | 6.684 | 6.6 | 6.7 | 7.2 | 7.5 | 6.1 | 7.8 | 5.9 | 6.8 | 7 | 6.3453 | 7.7 | 7 | 6.1 | 5.6 | 8 |  |  |  | 108
| 1 | 6 | 6.552 | 6.7 | 7.7 | 7 | 7.3 | 6.5 | 7.7 | 5.8 | 7 | 6.9 | 6.1851 | 7.5 | 7 | 6.7 | 5.5 | 8
| Mean |  | 6.696 | 6.65 | 7.0167 | 7.1 | 7.4833 | 6.6667 | 7.6 | 5.85 | 6.9333 | 6.7833 | 6.2017 | 7.59 | 7 | 6.7 | 5.7 | 7.3333
| Median |  | 6.648 | 6.65 | 7 | 7.1 | 7.5 | 6.75 | 7.7 | 5.85 | 6.95 | 6.75 | 6.2092 | 7.58 | 7 | 6.55 | 5.6 | 7
| Std Dev. |  | 0.21466 | 0.054772 | 0.47081 | 0.089443 | 0.31252 | 0.3266 | 0.21909 | 0.054772 | 0.12111 | 0.1472 | 0.16353 | 0.10621 | 0 | 0.7823 | 0.35214 | 0.5164
| Rel.Std.Dev. |  | 3.206% | 0.8236% | 6.71% | 1.26% | 4.176% | 4.899% | 2.883% | 0.9363% | 1.747% | 2.17% | 2.637% | 1.399% | 0% | 11.68% | 6.178% | 7.042%
| PDM3 |  | -1.288% | -1.966% | 3.439% | 4.668% | 10.32% | -1.72% | 12.04% | -13.76% | 2.211% | -0.0003946% | -8.575% | 11.89% | 3.194% | -1.229% | -15.97% | 8.108%
| Z-Score (Absolute) |  | 0.1 | 0.2 | 0.59 | 0.76 | 1.58 | 0.16 | 1.83 | 1.9 | 0.41 | 0.09 | 1.15 | 1.81 | NA | 0.09 | 2.22 | NA
Results from laboratories 2.15 and 2.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 290 (Certified Value 2.09 ppm) | [20 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [5 empty] | 2
| 1 | 1 | 3.0732 | 2.6 | 2.3 | 2 | 1.7 | 5 | 1 | 2 | 2.1678 | 2.09 | 2 | 1.4 | < 5 | 2.9 | [5 empty] | 1
| 1 | 2 | 2.6998 | 2.3 | 2.1 | 2 | 1.8 | 4.8 | 1 | 2 | 1.696 | 2 | 2 | 1.5 | < 5 | 3.01 | [5 empty] | 48
| 1 | 3 | 2.9334 | 2.4 | 1.9 | 2.4 | 2.1 | 4.8 | 1 | 2 | 1.6434 | 2.13 | 2 | 1.7 | < 5 | 2.84 | [5 empty] | 16
| 1 | 4 | 2.8704 | 2.4 | 1.7 | 2.2 | 1.9 | 5.5 | 1 | 2 | 1.9647 | 2 | 2 | 1.5 | < 5 | 2.43 | [5 empty] | 2.0852
| 1 | 5 | 2.8025 | 2.5 | 1.9 | 2 | 2 | 4.6 | 2 | 2 | 1.9142 | 2.1 | 2 | 1.4 | < 5 | 3.22 | [5 empty] | 109
| 1 | 6 | 2.4705 | 2.2 | 2.1 | 2 | 1.9 | 4.9 | 1 | 2 | 2.3053 | 1.99 | 2 | 1.7 | < 5 | 3.35
| Mean |  | 2.8083 | 2.4 | 2 | 2.1 | 1.9 | 4.9333 | 1.1667 | 2 | 1.9486 | 2.0517 | 2 | 1.5333 |  | 2.9583
| Median |  | 2.8365 | 2.4 | 2 | 2 | 1.9 | 4.85 | 1 | 2 | 1.9394 | 2.045 | 2 | 1.5 |  | 2.955
| Std Dev. |  | 0.20765 | 0.14142 | 0.20976 | 0.16733 | 0.14142 | 0.30768 | 0.40825 | 0 | 0.25834 | 0.061779 | 0 | 0.13663 |  | 0.32283
| Rel.Std.Dev. |  | 7.394% | 5.893% | 10.49% | 7.968% | 7.443% | 6.237% | 34.99% | 0% | 13.26% | 3.011% | 0% | 8.91% |  | 10.91%
| PDM3 |  | 34.68% | 15.1% | -4.087% | 0.7082% | -8.883% | 136.6% | -44.05% | -4.087% | -6.554% | -1.61% | -4.087% | -26.47% |  | 41.87%
| Z-Score (Absolute) |  | 1.59 | 0.67 | 0.22 | 0 | 0.45 | 6.37 | NA | NA | 0.34 | 0.11 | NA | 1.27 | 0.9 | 1.93
Results from laboratories 2.09, 2.11 and 2.15 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 290 (Indicative Value 3.85 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 3.986 | 3.8 | [17 empty] | 1
| 1 | 2 | 4.006 | 3.5 | [17 empty] | 18
| 1 | 3 | 4.001 | 3.3 | [17 empty] | 16
| 1 | 4 | 4.023 | 3.9 | [17 empty] | 3.8509
| 1 | 5 | 3.985 | 3.6 | [17 empty] | 24
| 1 | 6 | 4.01 | 4.1
| Mean |  | 4.0018 | 3.7
| Median |  | 4.0035 | 3.7
| Std Dev. |  | 0.014607 | 0.28983
| Rel.Std.Dev. |  | 0.365% | 7.833%
| PDM3 |  | 3.919% | -3.919%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 290 (Certified Value 1.56 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [5 empty] | 3
| 1 | 1 | 1.5 | 1.5 | 1.4 | 1.7 | 1.7 | < 10 | 1.6 | < 10 | 1.5853 | 1.66 | 1.6 | 1.56 | < 10 | 1.7 | [5 empty] | 1
| 1 | 2 | 1.4 | 1.5 | 1.4 | 1.6 | 1.8 | < 10 | 1.6 | < 10 | 1.6275 | 1.62 | 1.6 | 1.49 | < 10 | 1.8 | [5 empty] | 6
| 1 | 3 | 1.4 | 1.5 | 1.5 | 1.6 | 1.8 | < 10 | 1.5 | < 10 | 1.6194 | 1.62 | 1.5 | 1.62 | < 10 | 1.8 | [5 empty] | 16
| 1 | 4 | 1.3 | 1.5 | 1.4 | 1.4 | 1.8 | < 10 | 1.6 | < 10 | 1.6424 | 1.65 | 1.6 | 1.56 | < 10 | 1.7 | [5 empty] | 1.5634
| 1 | 5 | 1.4 | 1.6 | 1.4 | 1.4 | 1.8 | < 10 | 1.7 | < 10 | 1.6128 | 1.69 | 1.6 | 1.55 | < 10 | 1.7 | [5 empty] | 110
| 1 | 6 | 1.6 | 1.5 | 1.4 | 1.6 | 1.7 | < 10 | 1.7 | < 10 | 1.5996 | 1.67 | 1.5 | 1.53 | < 10 | 1.6
| Mean |  | 1.4333 | 1.5167 | 1.4167 | 1.55 | 1.7667 |  | 1.6167 |  | 1.6145 | 1.6517 | 1.5667 | 1.5517 |  | 1.7167
| Median |  | 1.4 | 1.5 | 1.4 | 1.6 | 1.8 |  | 1.6 |  | 1.6161 | 1.655 | 1.6 | 1.555 |  | 1.7
| Std Dev. |  | 0.10328 | 0.040825 | 0.040825 | 0.12247 | 0.05164 |  | 0.075277 |  | 0.020235 | 0.027869 | 0.05164 | 0.042622 |  | 0.075277
| Rel.Std.Dev. |  | 7.206% | 2.692% | 2.882% | 7.902% | 2.923% |  | 4.656% |  | 1.253% | 1.687% | 3.296% | 2.747% |  | 4.385%
| PDM3 |  | -8.322% | -2.992% | -9.388% | -0.8601% | 13% |  | 3.404% |  | 3.264% | 5.643% | 0.2059% | -0.7535% |  | 9.8%
| Z-Score (Absolute) |  | 1.23 | 0.67 | 1.34 | 0.44 | 1.02 | 22.82 | 0.01 | 22.82 | 0.01 | 0.24 | 0.33 | 0.43 | 22.82 | 0.68
| Analytical results for Sr in OREAS 290 (Certified Value 32.6 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g |  | 2
| 1 | 1 | 30.5401 | 30.2 | 33.9 | 33.7 | 31.4 | 34.4 | 33.4 | 28.7 | 31.2 | 32.4 | 33.7867 | 33.85 | 31.7 | 28 | 34.7 | 35 | 32.759 | 29.1 |  | 1
| 1 | 2 | 30.3603 | 30.5 | 32.6 | 32.7 | 34 | 32.2 | 33.8 | 29.7 | 31.4 | 32.3 | 34.9856 | 34.41 | 33.4 | 35 | 34 | 34 | 32.4728 | 31.1 |  | 7
| 1 | 3 | 29.6812 | 30.8 | 32.8 | 34.3 | 35.4 | 33.4 | 34.1 | 29.2 | 30.7 | 32.8 | 33.6134 | 33.69 | 32.6 | 31 | 33.1 | 36 | 31.7177 | 30.4 |  | 16
| 1 | 4 | 30.1574 | 30.8 | 31 | 33 | 33.4 | 29.9 | 34.4 | 29.5 | 30.3 | 32.8 | 33.3727 | 34.39 | 32.2 | 38 | 34.1 | 36 | 31.8495 | 33.6 |  | 32.6002
| 1 | 5 | 30.5382 | 30.8 | 32.5 | 34 | 33.1 | 28.8 | 34.3 | 29.1 | 30.9 | 33.6 | 34.0925 | 34.82 | 33.2 | 29 | 36.9 | 35 | 32.2873 | 36.1 |  | 111
| 1 | 6 | 30.4747 | 30.9 | 36.8 | 32.6 | 33 | 32.7 | 34.7 | 29 | 31.8 | 33.7 | 34.7565 | 34.71 | 31.7 | 32 | 35.6 | 34 | 32.9417 | 32.9
| Mean |  | 30.292 | 30.6667 | 33.2667 | 33.3833 | 33.3833 | 31.9 | 34.1167 | 29.2 | 31.05 | 32.9333 | 34.1012 | 34.3117 | 32.4667 | 32.1667 | 34.7333 | 35 | 32.338 | 32.2
| Median |  | 30.4175 | 30.8 | 32.7 | 33.35 | 33.25 | 32.45 | 34.2 | 29.15 | 31.05 | 32.8 | 33.9396 | 34.4 | 32.4 | 31.5 | 34.4 | 35 | 32.38 | 32
| Std Dev. |  | 0.33209 | 0.26583 | 1.9633 | 0.71391 | 1.3121 | 2.1373 | 0.46224 | 0.35777 | 0.53198 | 0.59217 | 0.64484 | 0.45442 | 0.73121 | 3.7639 | 1.3456 | 0.89443 | 0.487 | 2.5187
| Rel.Std.Dev. |  | 1.096% | 0.8668% | 5.902% | 2.139% | 3.93% | 6.7% | 1.355% | 1.225% | 1.713% | 1.798% | 1.891% | 1.324% | 2.252% | 11.7% | 3.874% | 2.556% | 1.506% | 7.822%
| PDM3 |  | -7.08% | -5.931% | 2.044% | 2.402% | 2.402% | -2.148% | 4.652% | -10.43% | -4.755% | 1.022% | 4.604% | 5.25% | -0.4095% | -1.33% | 6.543% | 7.361% | -0.8041% | -1.227%
| Z-Score (Absolute) |  | 1.48 | 1.25 | 0.35 | 0.42 | 0.42 | 0.49 | 0.87 | 2.14 | 1.01 | 0.14 | 0.86 | 0.99 | 0.14 | 0.33 | 1.25 | 1.41 | 0.22 | 0.31
| Analytical results for Ta in OREAS 290 (Certified Value < 0.05 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | [8 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 5 | [8 empty] | 1
| 1 | 2 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 5 | [8 empty] | 35
| 1 | 3 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 5 | [8 empty] | 16
| 1 | 4 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 5 | [8 empty] | < 0.05
| 1 | 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 5 | [8 empty] | 112
| 1 | 6 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.07 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 5
| Mean |  |  |  |  | 0.01 |  | 0.06
| Median |  |  |  |  | 0.01 |  | 0.06
| Std Dev. |  |  |  |  | 0 |  | 0.014142
| Rel.Std.Dev. |  |  |  |  | 0% |  | 23.57%
PDM3
| Z-Score (Absolute) |  | 0.77 | 0.77 | 0.77 | 0.67 | 0 | 0.45 | 0 | 0.77 | 0 | 0 | 95.37
| Analytical results for Tb in OREAS 290 (Indicative Value 0.37 ppm) | [20 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.09 | 2.11 | 2.14 | 2.17 | [15 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [15 empty] | 2
|  |  | 0.25g | 0.25g | 0.5g | 0.5g | [15 empty] | 2
| 1 | 1 | 0.36 | 0.38 | 0.399 | 0.4 | [15 empty] | 1
| 1 | 2 | 0.36 | 0.39 | 0.392 | 0.4 | [15 empty] | 19
| 1 | 3 | 0.37 | 0.36 | 0.385 | 0.3 | [15 empty] | 16
| 1 | 4 | 0.37 | 0.36 | 0.395 | 0.3 | [15 empty] | 0.37296
| 1 | 5 | 0.39 | 0.38 | 0.4 | 0.3 | [15 empty] | 25
| 1 | 6 | 0.38 | 0.39 | 0.39 | 0.4
| Mean |  | 0.37167 | 0.37667 | 0.3935 | 0.35
| Median |  | 0.37 | 0.38 | 0.3935 | 0.35
| Std Dev. |  | 0.01169 | 0.013663 | 0.0056833 | 0.054772
| Rel.Std.Dev. |  | 3.145% | 3.627% | 1.444% | 15.65%
| PDM3 |  | -0.3463% | 0.9943% | 5.508% | -6.156%
| Z-Score (Absolute) |  | 0.15 | 0.15 | 1.19 | 1.49
| Analytical results for Te in OREAS 290 (Certified Value 0.053 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.1g | [5 empty] | 3
| 1 | 1 | 0.07 | 0.06 | 0.05 | 0.06 | 0.13 | < 0.05 | 0.05 | 0.059365 | 0.05 | < 0.05 | < 2 | 0.04 | < 5 | 0.05 | [5 empty] | 1
| 1 | 2 | 0.07 | 0.05 | 0.06 | 0.04 | 0.12 | 0.07 | < 0.05 | 0.05294 | 0.05 | 0.05 | < 2 | < 0.02 | < 5 | 0.05 | [5 empty] | 37
| 1 | 3 | 0.09 | 0.04 | 0.06 | 0.06 | 0.07 | < 0.05 | < 0.05 | 0.060545 | 0.06 | < 0.05 | < 2 | < 0.02 | < 5 | 0.05 | [5 empty] | 16
| 1 | 4 | 0.04 | 0.04 | 0.05 | 0.05 | 0.03 | < 0.05 | 0.05 | 0.053429 | 0.05 | < 0.05 | < 2 | 0.02 | < 5 | 0.04 | [5 empty] | 0.053157
| 1 | 5 | 0.05 | 0.04 | 0.06 | 0.05 | 0.05 | 0.06 | 0.05 | < 0.05 | 0.06 | < 0.05 | < 2 | < 0.02 | < 5 | 0.04 | [5 empty] | 113
| 1 | 6 | 0.07 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | < 0.05 | 0.05 | 0.05 | < 2 | 0.02 | < 5 | 0.05
| Mean |  | 0.065 | 0.046667 | 0.055 | 0.053333 | 0.075 | 0.06 | 0.05 | 0.05657 | 0.053333 | 0.05 |  | 0.026667 |  | 0.046667
| Median |  | 0.07 | 0.045 | 0.055 | 0.055 | 0.06 | 0.06 | 0.05 | 0.056397 | 0.05 | 0.05 |  | 0.02 |  | 0.05
| Std Dev. |  | 0.017607 | 0.008165 | 0.0054772 | 0.008165 | 0.040866 | 0.01 | 0 | 0.0039435 | 0.005164 | 0 |  | 0.011547 |  | 0.005164
| Rel.Std.Dev. |  | 27.09% | 17.5% | 9.959% | 15.31% | 54.49% | 16.67% | 0% | 6.971% | 9.682% | 0% |  | 43.3% |  | 11.07%
| PDM3 |  | 22.28% | -12.21% | 3.467% | 0.3318% | 41.09% | 12.87% | -5.939% | 6.42% | 0.3318% | -5.939% |  | -49.83% |  | -12.21%
| Z-Score (Absolute) |  | 1.28 | 0.28 | 0.43 | 0.28 | 2.13 | 0.64 | 0.71 | 0.34 | 0.28 | 1.42 | 80.92 | 2.7 | 208.68 | 0.28
| Analytical results for Th in OREAS 290 (Certified Value 11.3 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.25g | 0.1g |  |  |  |  | 1
| 1 | 1 | 9.2045 | 11.56 | 9.4 | 11.6 | 11.3 | 11.3 | 11.1 | 10.8 | 9.6 | 12.7053 | 11.494 | 11.7 | 0.3 | < 20 | 13.2 |  |  |  |  | 1
| 1 | 2 | 8.3061 | 11.64 | 10.2 | 11 | 11.7 | 11.8 | 10.1 | 10.7 | 9.7 | 13.1473 | 11.482 | 11.7 | 0.3 | < 20 | 14 |  |  |  |  | 38
| 1 | 3 | 8.4734 | 12.13 | 9 | 11.3 | 11.6 | 12.1 | 10.3 | 11.6 | 9 | 12.8835 | 11.313 | 11.7 | 0.5 | < 20 | 13.6 |  |  |  |  | 16
| 1 | 4 | 7.9897 | 10.82 | 9.9 | 10.6 | 11.7 | 11.7 | 10.4 | 11.8 | 9.1 | 12.8265 | 11.688 | 11.5 | 0.3 | < 20 | 13.5 |  |  |  |  | 11.3145
| 1 | 5 | 7.7368 | 11.38 | 9.4 | 11.3 | 11.7 | 11.4 | 9.7 | 11.2 | 9.5 | 12.8082 | 11.4 | 11.6 | 0.3 | < 20 | 13.5 |  |  |  |  | 114
| 1 | 6 | 8.0606 | 11.26 | 9.5 | 12.8 | 11.3 | 11.5 | 10.7 | 10.9 | 9.4 | 13.2036 | 11.443 | 11.5 | 0.2 | < 20 | 12.9
| Mean |  | 8.2952 | 11.465 | 9.5667 | 11.4333 | 11.55 | 11.6333 | 10.3833 | 11.1667 | 9.3833 | 12.9291 | 11.47 | 11.6167 | 0.31667 |  | 13.45
| Median |  | 8.1834 | 11.47 | 9.45 | 11.3 | 11.65 | 11.6 | 10.35 | 11.05 | 9.45 | 12.855 | 11.4625 | 11.65 | 0.3 |  | 13.5
| Std Dev. |  | 0.51359 | 0.43524 | 0.42269 | 0.75011 | 0.19748 | 0.29439 | 0.48339 | 0.45019 | 0.27869 | 0.20016 | 0.12532 | 0.098319 | 0.098319 |  | 0.37283
| Rel.Std.Dev. |  | 6.191% | 3.796% | 4.418% | 6.561% | 1.71% | 2.531% | 4.655% | 4.032% | 2.97% | 1.548% | 1.093% | 0.8464% | 31.05% |  | 2.772%
| PDM3 |  | -26.69% | 1.33% | -15.45% | 1.05% | 2.081% | 2.818% | -8.23% | -1.307% | -17.07% | 14.27% | 1.374% | 2.671% | -97.2% |  | 18.87%
| Z-Score (Absolute) |  | 2.02 | 0.02 | 1.2 | 0 | 0.07 | 0.13 | 0.67 | 0.17 | 1.32 | 0.96 | 0.02 | 0.12 | 7.14 | 0.92 | 1.3
| Analytical results for Ti in OREAS 290 (Certified Value 0.119 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 3
| 1 | 1 | 0.11685 | 0.11156 | 0.096 | 0.126 | 0.119 | 0.125 | 0.118 | 0.12 | 0.11 | 0.12 | 0.11 | 0.15707 | 0.118 | 0.141 | 0.1031 | 0.11 | 0.216 | 0.12155 | 0.112 | 1
| 1 | 2 | 0.12015 | 0.11441 | 0.097 | 0.122 | 0.123 | 0.126 | 0.116 | 0.13 | 0.12 | 0.12 | 0.11 | 0.16282 | 0.1191 | 0.139 | 0.1288 | 0.1 | 0.18 | 0.1287 | 0.114 | 8
| 1 | 3 | 0.1323 | 0.11262 | 0.098 | 0.125 | 0.122 | 0.125 | 0.119 | 0.13 | 0.11 | 0.12 | 0.11 | 0.15826 | 0.1158 | 0.138 | 0.1134 | 0.12 | 0.483 | 0.12155 | 0.12 | 16
| 1 | 4 | 0.11985 | 0.11136 | 0.098 | 0.126 | 0.123 | 0.126 | 0.13 | 0.13 | 0.12 | 0.12 | 0.11 | 0.15526 | 0.117 | 0.139 | 0.1327 | 0.11 | 0.505 | 0.12155 | 0.131 | 0.11925
| 1 | 5 | 0.11805 | 0.1152 | 0.098 | 0.128 | 0.122 | 0.124 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.15968 | 0.1188 | 0.141 | 0.1049 | 0.1 | 0.456 | 0.10898 | 0.138 | 115
| 1 | 6 | 0.11415 | 0.11647 | 0.098 | 0.128 | 0.12 | 0.126 | 0.117 | 0.13 | 0.12 | 0.13 | 0.12 | 0.16204 | 0.1179 | 0.139 | 0.1138 | 0.1 | 0.359 | 0.12155 | 0.124
| Mean |  | 0.12023 | 0.1136 | 0.0975 | 0.12583 | 0.1215 | 0.12533 | 0.12 | 0.12833 | 0.11667 | 0.12167 | 0.11333 | 0.15919 | 0.11777 | 0.1395 | 0.11612 | 0.10667 | 0.3665 | 0.12065 | 0.12317
| Median |  | 0.11895 | 0.11351 | 0.098 | 0.126 | 0.122 | 0.1255 | 0.1185 | 0.13 | 0.12 | 0.12 | 0.11 | 0.15897 | 0.11795 | 0.139 | 0.1136 | 0.105 | 0.4075 | 0.12155 | 0.122
| Std Dev. |  | 0.0063084 | 0.0020782 | 0.00083666 | 0.0022286 | 0.0016432 | 0.0008165 | 0.005099 | 0.0040825 | 0.005164 | 0.0040825 | 0.005164 | 0.0029118 | 0.0012144 | 0.0012247 | 0.012197 | 0.008165 | 0.14018 | 0.0063899 | 0.010008
| Rel.Std.Dev. |  | 5.247% | 1.829% | 0.8581% | 1.771% | 1.352% | 0.6515% | 4.249% | 3.181% | 4.426% | 3.355% | 4.556% | 1.829% | 1.031% | 0.878% | 10.5% | 7.655% | 38.25% | 5.296% | 8.126%
| PDM3 |  | 0.8162% | -4.737% | -18.24% | 5.519% | 1.885% | 5.1% | 0.6275% | 7.615% | -2.168% | 2.025% | -4.963% | 33.49% | -1.245% | 16.98% | -2.629% | -10.55% | 207.3% | 1.17% | 3.283%
| Z-Score (Absolute) |  | 0.06 | 1.01 | 3.33 | 0.75 | 0.12 | 0.67 | 0.09 | 1.11 | 0.57 | 0.15 | 1.05 | 5.55 | 0.41 | 2.71 | 0.65 | 2.01 | 35.38 | 0 | 0.36
| Analytical results for Tl in OREAS 290 (Certified Value 0.52 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.1g |  |  |  |  | 3
| 1 | 1 | 0.5782 | 0.48 | 0.48 | 0.49 | 0.5 | 0.51 | 0.51 | 0.47 | 0.51574 | 0.55 | 0.55 | < 5 | 0.47 | < 10 | 0.55 |  |  |  |  | 1
| 1 | 2 | 0.5349 | 0.54 | 0.47 | 0.5 | 0.52 | 0.48 | 0.52 | 0.49 | 0.53496 | 0.5 | 0.52 | < 5 | 0.47 | < 10 | 0.58 |  |  |  |  | 40
| 1 | 3 | 0.5543 | 0.49 | 0.5 | 0.48 | 0.55 | 0.52 | 0.54 | 0.46 | 0.51752 | 0.52 | 0.52 | < 5 | 0.51 | < 10 | 0.58 |  |  |  |  | 16
| 1 | 4 | 0.5369 | 0.55 | 0.47 | 0.51 | 0.53 | 0.39 | 0.53 | 0.46 | 0.51744 | 0.53 | 0.52 | < 5 | 0.5 | < 10 | 0.57 |  |  |  |  | 0.51524
| 1 | 5 | 0.5271 | 0.46 | 0.5 | 0.51 | 0.53 | 0.47 | 0.53 | 0.48 | 0.51153 | 0.54 | 0.53 | < 5 | 0.49 | < 10 | 0.56 |  |  |  |  | 116
| 1 | 6 | 0.547 | 0.49 | 0.56 | 0.51 | 0.52 | 0.51 | 0.51 | 0.48 | 0.53527 | 0.53 | 0.53 | < 5 | 0.49 | < 10 | 0.56
| Mean |  | 0.5464 | 0.50167 | 0.49667 | 0.5 | 0.525 | 0.48 | 0.52333 | 0.47333 | 0.52207 | 0.52833 | 0.52833 |  | 0.48833 |  | 0.56667
| Median |  | 0.54195 | 0.49 | 0.49 | 0.505 | 0.525 | 0.495 | 0.525 | 0.475 | 0.51748 | 0.53 | 0.525 |  | 0.49 |  | 0.565
| Std Dev. |  | 0.01827 | 0.035449 | 0.033862 | 0.012649 | 0.016432 | 0.048166 | 0.012111 | 0.012111 | 0.010331 | 0.017224 | 0.01169 |  | 0.016021 |  | 0.012111
| Rel.Std.Dev. |  | 3.344% | 7.066% | 6.818% | 2.53% | 3.13% | 10.03% | 2.314% | 2.559% | 1.979% | 3.26% | 2.213% |  | 3.281% |  | 2.137%
| PDM3 |  | 6.047% | -2.635% | -3.605% | -2.958% | 1.894% | -6.84% | 1.57% | -8.134% | 1.326% | 2.541% | 2.541% |  | -5.222% |  | 9.981%
| Z-Score (Absolute) |  | 0.67 | 0.63 | 0.77 | 0.67 | 0.05 | 1.25 | 0 | 1.44 | 0.04 | 0.14 | 0.14 | 57.12 | 1.01 | 129.37 | 1.25
| Analytical results for Tm in OREAS 290 (Indicative Value 0.11 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 2.14 | 2.17 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.12 | 0.1 | [17 empty] | 1
| 1 | 2 | 0.119 | 0.1 | [17 empty] | 20
| 1 | 3 | 0.116 | 0.1 | [17 empty] | 16
| 1 | 4 | 0.12 | 0.1 | [17 empty] | 0.10983
| 1 | 5 | 0.123 | 0.1 | [17 empty] | 26
| 1 | 6 | 0.12 | 0.1
| Mean |  | 0.11967 | 0.1
| Median |  | 0.12 | 0.1
| Std Dev. |  | 0.0022509 | 1.5202e-17
| Rel.Std.Dev. |  | 1.881% | 1.52e-14%
| PDM3 |  | 8.953% | -8.953%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 290 (Certified Value 1.38 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.25g | 0.1g | [5 empty] | 2
| 1 | 1 | 1.1885 | 1.13 | 1.32 | 1.35 | 1.45 | 1.33 | 1.37 | 1.15 | 1.5955 | 1.469 | 1.7 | 1.4 | 2.1 | 1.53 | [5 empty] | 1
| 1 | 2 | 1.0871 | 1.3 | 1.27 | 1.37 | 1.53 | 1.27 | 1.36 | 1.16 | 1.5772 | 1.45 | 1.7 | 1.4 | 1.6 | 1.55 | [5 empty] | 42
| 1 | 3 | 1.1141 | 1.23 | 1.35 | 1.39 | 1.56 | 1.25 | 1.47 | 1.14 | 1.6005 | 1.404 | 1.71 | 1.5 | 2.3 | 1.57 | [5 empty] | 16
| 1 | 4 | 1.0709 | 1.26 | 1.26 | 1.4 | 1.5 | 1.22 | 1.44 | 1.16 | 1.5755 | 1.464 | 1.64 | 1.4 | 2.7 | 1.57 | [5 empty] | 1.3848
| 1 | 5 | 1.0433 | 1.2 | 1.34 | 1.4 | 1.44 | 1.19 | 1.45 | 1.13 | 1.6173 | 1.407 | 1.66 | 1.4 | 3.2 | 1.56 | [5 empty] | 117
| 1 | 6 | 1.0763 | 1.26 | 1.51 | 1.34 | 1.47 | 1.32 | 1.41 | 1.16 | 1.6278 | 1.44 | 1.66 | 1.4 | 2.8 | 1.39
| Mean |  | 1.0967 | 1.23 | 1.3417 | 1.375 | 1.4917 | 1.2633 | 1.4167 | 1.15 | 1.599 | 1.439 | 1.6783 | 1.4167 | 2.45 | 1.5283
| Median |  | 1.0817 | 1.245 | 1.33 | 1.38 | 1.485 | 1.26 | 1.425 | 1.155 | 1.598 | 1.445 | 1.68 | 1.4 | 2.5 | 1.555
| Std Dev. |  | 0.050508 | 0.05933 | 0.090203 | 0.025884 | 0.047081 | 0.055015 | 0.044572 | 0.012649 | 0.021009 | 0.027914 | 0.028577 | 0.040825 | 0.56833 | 0.069402
| Rel.Std.Dev. |  | 4.605% | 4.824% | 6.723% | 1.882% | 3.156% | 4.355% | 3.146% | 1.1% | 1.314% | 1.94% | 1.703% | 2.882% | 23.2% | 4.541%
| PDM3 |  | -20.8% | -11.18% | -3.113% | -0.7054% | 7.72% | -8.769% | 2.304% | -16.95% | 15.47% | 3.916% | 21.2% | 2.304% | 76.92% | 10.37%
| Z-Score (Absolute) |  | 1.63 | 0.95 | 0.38 | 0.21 | 0.38 | 0.78 | 0 | 1.36 | 0.93 | 0.11 | 1.33 | 0 | 5.26 | 0.57
| Analytical results for V in OREAS 290 (Certified Value 61 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 1
| 1 | 1 | 54.495 | 58.03 | 62 | 62 | 61 | 61 | 61 | 62 | 52 | 63 | 62 | 66.0911 | 67.7 | 58 | 51 | 59 | 26.9 | 62.1 | 67 | 1
| 1 | 2 | 55.02 | 59.03 | 62 | 60 | 61 | 61 | 61 | 63 | 54 | 63 | 60 | 65.8161 | 69.3 | 57 | 64 | 60 | 23.8 | 62.69 | 69 | 9
| 1 | 3 | 59.115 | 58.43 | 63 | 61 | 61 | 62 | 62 | 64 | 53 | 60 | 60 | 62.9629 | 68.2 | 56 | 56 | 64 | 40.2 | 61.46 | 66 | 16
| 1 | 4 | 54.81 | 59.13 | 64 | 62 | 62 | 61 | 65 | 64 | 54 | 62 | 62 | 63.1859 | 69.1 | 57 | 70 | 62 | 41.6 | 61.7108 | 78 | 60.7722
| 1 | 5 | 54.705 | 59.19 | 63 | 63 | 62 | 60 | 62 | 64 | 53 | 62 | 64 | 62.5529 | 70.8 | 58 | 53 | 61 | 39.3 | 61.615 | 84 | 118
| 1 | 6 | 54.18 | 59.8 | 64 | 62 | 61 | 61 | 61 | 64 | 53 | 63 | 62 | 63.362 | 68.7 | 57 | 58 | 63 | 32.9 | 61.585 | 67
| Mean |  | 55.3875 | 58.935 | 63 | 61.6667 | 61.3333 | 61 | 62 | 63.5 | 53.1667 | 62.1667 | 61.6667 | 63.9952 | 68.9667 | 57.1667 | 58.6667 | 61.5 | 34.1167 | 61.8601 | 71.8333
| Median |  | 54.7575 | 59.08 | 63 | 62 | 61 | 61 | 61.5 | 64 | 53 | 62.5 | 62 | 63.274 | 68.9 | 57 | 57 | 61.5 | 36.1 | 61.6629 | 68
| Std Dev. |  | 1.8484 | 0.62199 | 0.89443 | 1.0328 | 0.5164 | 0.63246 | 1.5492 | 0.83666 | 0.75277 | 1.169 | 1.5055 | 1.5434 | 1.0727 | 0.75277 | 7.1461 | 1.8708 | 7.4816 | 0.46139 | 7.414
| Rel.Std.Dev. |  | 3.337% | 1.055% | 1.42% | 1.675% | 0.842% | 1.037% | 2.499% | 1.318% | 1.416% | 1.881% | 2.441% | 2.412% | 1.555% | 1.317% | 12.18% | 3.042% | 21.93% | 0.7459% | 10.32%
| PDM3 |  | -8.86% | -3.023% | 3.666% | 1.472% | 0.9234% | 0.3749% | 2.02% | 4.489% | -12.51% | 2.295% | 1.472% | 5.303% | 13.48% | -5.933% | -3.465% | 1.198% | -43.86% | 1.79% | 18.2%
| Z-Score (Absolute) |  | 2.31 | 1 | 0.49 | 0 | 0.12 | 0.25 | 0.12 | 0.67 | 3.13 | 0.18 | 0 | 0.86 | 2.68 | 1.66 | 1.1 | 0.06 | 10.13 | 0.07 | 3.74
| Analytical results for W in OREAS 290 (Certified Value 0.54 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.1g |  |  | 2
| 1 | 1 | 0.4651 | 0.48 | 0.49 | 0.53 | 0.5 | 0.55 | 0.7 | < 10 | 0.6 | < 10 | 0.60361 | 0.49 | 0.4 | < 2 | < 0.1 | < 10 | 0.69 |  |  | 1
| 1 | 2 | 0.4868 | 0.53 | 0.47 | 0.52 | 0.48 | 0.5 | 0.5 | < 10 | 0.6 | < 10 | 0.57809 | 0.71 | 0.4 | < 2 | < 0.1 | < 10 | 0.68 |  |  | 10
| 1 | 3 | 0.4764 | 0.49 | 0.47 | 0.52 | 0.55 | 0.57 | 0.5 | < 10 | 0.6 | < 10 | 0.60273 | 0.54 | 0.4 | < 2 | < 0.1 | < 10 | 0.73 |  |  | 16
| 1 | 4 | 0.4893 | 0.54 | 0.41 | 0.55 | 0.51 | 0.45 | 0.6 | < 10 | 0.6 | < 10 | 0.6179 | 0.59 | 0.4 | < 2 | < 0.1 | < 10 | 0.7 |  |  | 0.54297
| 1 | 5 | 0.4793 | 0.53 | 0.45 | 0.53 | 0.53 | 0.39 | 0.5 | 12 | 0.6 | < 10 | 0.6199 | 0.6 | 0.4 | < 2 | < 0.1 | < 10 | 0.71 |  |  | 119
| 1 | 6 | 0.461 | 0.51 | 0.52 | 0.55 | 0.52 | 0.45 | 0.5 | < 10 | 0.6 | < 10 | 0.57622 | 0.59 | 0.4 | < 2 | < 0.1 | < 10 | 0.6
| Mean |  | 0.47632 | 0.51333 | 0.46833 | 0.53333 | 0.515 | 0.485 | 0.55 | 12 | 0.6 |  | 0.59974 | 0.58667 | 0.4 |  |  |  | 0.685
| Median |  | 0.47785 | 0.52 | 0.47 | 0.53 | 0.515 | 0.475 | 0.5 | 12 | 0.6 |  | 0.60317 | 0.59 | 0.4 |  |  |  | 0.695
| Std Dev. |  | 0.011382 | 0.024221 | 0.037103 | 0.013663 | 0.02429 | 0.068044 | 0.083666 |  | 0 |  | 0.018876 | 0.073394 | 6.0809e-17 |  |  |  | 0.045056
| Rel.Std.Dev. |  | 2.39% | 4.718% | 7.922% | 2.562% | 4.716% | 14.03% | 15.21% |  | 0% |  | 3.147% | 12.51% | 1.52e-14% |  |  |  | 6.577%
| PDM3 |  | -12.28% | -5.459% | -13.75% | -1.775% | -5.152% | -10.68% | 1.294% | 2110% | 10.5% |  | 10.46% | 8.047% | -26.33% |  |  |  | 26.16%
| Z-Score (Absolute) |  | 0.67 | 0.34 | 0.75 | 0.15 | 0.32 | 0.59 | 0 | 51.4 | NA | 40.72 | 0.46 | 0.34 | NA | 4.12 | 4.58 | 40.72 | 1.24
Results from laboratories 2.11 and 2.15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 290 (Certified Value 8.21 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [5 empty] | 3
| 1 | 1 | 8.29 | 8.46 | 8.06 | 8.16 | 7.94 | 7 | 7.51 | 8.2 | 9.7351 | 8.77 | 8.56 | 7.84 | 7.5285 | 8.87 | [5 empty] | 1
| 1 | 2 | 7.95 | 8.22 | 8.58 | 7.67 | 7.91 | 7 | 7.5 | 8.1 | 9.9669 | 8.5 | 8.77 | 7.86 | 7.605 | 8.97 | [5 empty] | 11
| 1 | 3 | 8.2 | 8.57 | 8.93 | 7.78 | 8.24 | 7 | 7.43 | 8.1 | 9.6537 | 8.68 | 8.64 | 7.77 | 7.371 | 9.08 | [5 empty] | 16
| 1 | 4 | 7.61 | 8.37 | 8.44 | 7.71 | 8.42 | 7 | 7.47 | 8.3 | 9.5707 | 8.74 | 8.63 | 7.92 | 7.371 | 9.34 | [5 empty] | 8.2067
| 1 | 5 | 8.13 | 8.44 | 8.56 | 7.08 | 8.5 | 7 | 7.43 | 8.9 | 9.8 | 8.87 | 8.76 | 8.14 | 7.5375 | 8.96 | [5 empty] | 120
| 1 | 6 | 9.28 | 8.3 | 8.37 | 7.56 | 8.16 | 7 | 7.5 | 8.7 | 9.8842 | 8.9 | 8.36 | 8.16 | 7.5465 | 8.26
| Mean |  | 8.2433 | 8.3933 | 8.49 | 7.66 | 8.195 | 7 | 7.4733 | 8.3833 | 9.7684 | 8.7433 | 8.62 | 7.9483 | 7.4933 | 8.9133
| Median |  | 8.165 | 8.405 | 8.5 | 7.69 | 8.2 | 7 | 7.485 | 8.25 | 9.7676 | 8.755 | 8.635 | 7.89 | 7.533 | 8.965
| Std Dev. |  | 0.56191 | 0.1242 | 0.2858 | 0.35026 | 0.24214 | 0 | 0.036148 | 0.33714 | 0.14642 | 0.14459 | 0.15086 | 0.16351 | 0.098402 | 0.35887
| Rel.Std.Dev. |  | 6.817% | 1.48% | 3.366% | 4.573% | 2.955% | 0% | 0.4837% | 4.022% | 1.499% | 1.654% | 1.75% | 2.057% | 1.313% | 4.026%
| PDM3 |  | 0.4469% | 2.275% | 3.453% | -6.661% | -0.1421% | -14.7% | -8.936% | 2.153% | 19.03% | 6.539% | 5.037% | -3.148% | -8.693% | 8.611%
| Z-Score (Absolute) |  | 0.26 | 0.02 | 0.2 | 1.35 | 0.35 | NA | 1.7 | 0 | 2.59 | 0.67 | 0.44 | 0.81 | 1.67 | 0.99
Results from laboratory 2.10 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 290 (Indicative Value 0.77 ppm) | [20 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.09 | 2.11 | 2.14 | 2.17 | [15 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [15 empty] | 2
|  |  | 0.25g | 0.25g | 0.5g | 0.5g | [15 empty] | 2
| 1 | 1 | 0.7 | 0.7 | 0.751 | 0.8 | [15 empty] | 1
| 1 | 2 | 0.7 | 0.8 | 0.769 | 0.8 | [15 empty] | 21
| 1 | 3 | 0.8 | 0.8 | 0.719 | 0.8 | [15 empty] | 16
| 1 | 4 | 0.8 | 0.8 | 0.759 | 0.8 | [15 empty] | 0.77363
| 1 | 5 | 0.8 | 0.8 | 0.762 | 0.8 | [15 empty] | 27
| 1 | 6 | 0.7 | 0.8 | 0.765 | 0.8
| Mean |  | 0.75 | 0.78333 | 0.75417 | 0.8
| Median |  | 0.75 | 0.8 | 0.7605 | 0.8
| Std Dev. |  | 0.054772 | 0.040825 | 0.018269 | 1.2162e-16
| Rel.Std.Dev. |  | 7.303% | 5.212% | 2.422% | 1.52e-14%
| PDM3 |  | -3.055% | 1.254% | -2.516% | 3.408%
| Z-Score (Absolute) |  | 0.76 | 0.59 | 0.59 | 1.26
| Analytical results for Zn in OREAS 290 (Certified Value 86 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.1g | 1
| 1 | 1 | 91.2507 | 33.48 | 84.2 | 86 | 85 | 86 | 86 | 84 | 76 | 86 | 83 | 90.1655 | 89.1 | 95 | 81 | 89.2 | 60 | 90.8406 | 88.7 | 1
| 1 | 2 | 83.3265 | 32.7 | 82.2 | 83 | 85 | 86 | 85 | 84 | 74 | 84 | 84 | 90.6275 | 91.8 | 93 | 98 | 86 | 58 | 90.1476 | 87.4 | 12
| 1 | 3 | 87.8852 | 31.92 | 82.7 | 85 | 85 | 87 | 87 | 85 | 73 | 81 | 84 | 86.3269 | 89.8 | 92 | 81 | 90.5 | 68 | 90.1962 | 87.9 | 16
| 1 | 4 | 84.2181 | 30.81 | 84 | 86 | 87 | 87 | 88 | 86 | 73 | 82 | 86 | 86.2089 | 90.8 | 95 | 98 | 88.7 | 68 | 90.4986 | 88.6 | 86.2117
| 1 | 5 | 82.8432 | 31.47 | 83.3 | 86 | 86 | 85 | 87 | 86 | 71 | 83 | 86 | 87.701 | 91.4 | 95 | 81 | 90.5 | 71 | 90.6444 | 89.7 | 121
| 1 | 6 | 84.2147 | 31.31 | 85.3 | 85 | 84 | 87 | 86 | 85 | 73 | 85 | 85 | 87.0721 | 91.3 | 93 | 82 | 93.2 | 63 | 90.0396 | 88.8
| Mean |  | 85.6231 | 31.9483 | 83.6167 | 85.1667 | 85.3333 | 86.3333 | 86.5 | 85 | 73.3333 | 83.5 | 84.6667 | 88.017 | 90.7 | 93.8333 | 86.8333 | 89.6833 | 64.6667 | 90.3945 | 88.5167
| Median |  | 84.2164 | 31.695 | 83.65 | 85.5 | 85 | 86.5 | 86.5 | 85 | 73 | 83.5 | 84.5 | 87.3865 | 91.05 | 94 | 81.5 | 89.85 | 65.5 | 90.3474 | 88.65
| Std Dev. |  | 3.2786 | 0.98426 | 1.1197 | 1.169 | 1.0328 | 0.8165 | 1.0488 | 0.89443 | 1.633 | 1.8708 | 1.2111 | 1.9261 | 1.0431 | 1.3292 | 8.6583 | 2.3861 | 5.1251 | 0.31576 | 0.79352
| Rel.Std.Dev. |  | 3.829% | 3.081% | 1.339% | 1.373% | 1.21% | 0.9457% | 1.212% | 1.052% | 2.227% | 2.241% | 1.43% | 2.188% | 1.15% | 1.417% | 9.971% | 2.661% | 7.925% | 0.3493% | 0.8965%
| PDM3 |  | -0.6828% | -62.94% | -3.01% | -1.212% | -1.019% | 0.1411% | 0.3344% | -1.406% | -14.94% | -3.145% | -1.792% | 2.094% | 5.206% | 8.841% | 0.721% | 4.027% | -24.99% | 4.852% | 2.674%
| Z-Score (Absolute) |  | 0 | 17.05 | 0.64 | 0.14 | 0.09 | 0.23 | 0.28 | 0.2 | 3.9 | 0.67 | 0.3 | 0.76 | 1.61 | 2.61 | 0.38 | 1.29 | 6.66 | 1.52 | 0.92
| Analytical results for Zr in OREAS 290 (Certified Value 19.8 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [5 empty] | 1
| 1 | 1 | 17.9 | 17.8 | 17.1 | 16.7 | 17 | 9.6 | 16.9 | 20.1 | 23.9426 | 20.64 | 23.7 | 0.5 | 21.7249 | 20.5 | [5 empty] | 1
| 1 | 2 | 17.1 | 17.7 | 18.9 | 15.8 | 17.7 | 10 | 16.9 | 20.1 | 25.108 | 20.55 | 26.6 | 0.4 | 21.8317 | 18.5 | [5 empty] | 48
| 1 | 3 | 17.4 | 19.1 | 19.6 | 15.6 | 18 | 9.7 | 15.8 | 20.3 | 25.34 | 19.82 | 25.2 | 0.5 | 21.271 | 22.3 | [5 empty] | 16
| 1 | 4 | 15.9 | 18.3 | 18.1 | 16.9 | 18.9 | 10.1 | 16.2 | 21.1 | 24.2264 | 21.03 | 25 | 0.3 | 21.2131 | 22.7 | [5 empty] | 19.8408
| 1 | 5 | 17 | 18.5 | 18.2 | 14.8 | 19.1 | 10.4 | 16.8 | 22.5 | 24.7327 | 20.49 | 25.7 | 0.3 | 21.6181 | 20.8 | [5 empty] | 122
| 1 | 6 | 19.7 | 18 | 18.3 | 14.6 | 18 | 9.9 | 17.2 | 22.1 | 25.5151 | 20.34 | 24.8 | 0.2 | 21.9474 | 17.7
| Mean |  | 17.5 | 18.2333 | 18.3667 | 15.7333 | 18.1167 | 9.95 | 16.6333 | 21.0333 | 24.8108 | 20.4783 | 25.1667 | 0.36667 | 21.601 | 20.4167
| Median |  | 17.25 | 18.15 | 18.25 | 15.7 | 18 | 9.95 | 16.85 | 20.7 | 24.9203 | 20.52 | 25.1 | 0.35 | 21.6715 | 20.65
| Std Dev. |  | 1.2633 | 0.52026 | 0.83825 | 0.94587 | 0.77825 | 0.2881 | 0.52409 | 1.0558 | 0.62701 | 0.39706 | 0.96471 | 0.12111 | 0.29939 | 1.9984
| Rel.Std.Dev. |  | 7.219% | 2.853% | 4.564% | 6.012% | 4.296% | 2.895% | 3.151% | 5.02% | 2.527% | 1.939% | 3.833% | 33.03% | 1.386% | 9.788%
| PDM3 |  | -11.8% | -8.102% | -7.43% | -20.7% | -8.69% | -49.85% | -16.17% | 6.01% | 25.05% | 3.213% | 26.84% | -98.15% | 8.872% | 2.902%
| Z-Score (Absolute) |  | 0.23 | 0.02 | 0.02 | 0.73 | 0.05 | 2.37 | 0.47 | 0.78 | 1.85 | 0.62 | 1.95 | 5.1 | 0.94 | 0.6
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C558: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C576: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C612: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C649: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C667: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C722: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C758: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C776: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C812: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C830: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C904: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C922: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C940: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C958: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C976: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C994: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1012: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1048: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1066: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1084: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1122: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1140: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1158: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 290 (Indicative Value 14.09 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.062 | 1
| 1 | 2 | 14.119 | 23
| Mean |  | 14.0905 | 16
| Median |  | 14.0905 | 14.0905
| Std Dev. |  | 0.040305 | 29
| Rel.Std.Dev. |  | 0.286%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 290 (Indicative Value 0.742 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.752 | 1
| 1 | 2 | 0.731 | 24
| Mean |  | 0.7415 | 16
| Median |  | 0.7415 | 0.7415
| Std Dev. |  | 0.014849 | 30
| Rel.Std.Dev. |  | 2.003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 290 (Indicative Value 5.58 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.59 | 1
| 1 | 2 | 5.56 | 25
| Mean |  | 5.575 | 16
| Median |  | 5.575 | 5.575
| Std Dev. |  | 0.021213 | 31
| Rel.Std.Dev. |  | 0.3805%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 290 (Indicative Value 3.44 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.435 | 1
| 1 | 2 | 3.452 | 26
| Mean |  | 3.4435 | 16
| Median |  | 3.4435 | 3.4435
| Std Dev. |  | 0.012021 | 32
| Rel.Std.Dev. |  | 0.3491%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 290 (Indicative Value 2.36 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.32 | 1
| 1 | 2 | 2.39 | 23
| Mean |  | 2.355 | 16
| Median |  | 2.355 | 2.355
| Std Dev. |  | 0.049497 | 29
| Rel.Std.Dev. |  | 2.102%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 290 (Indicative Value 0.076 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.076 | 1
| 1 | 2 | 0.075 | 24
| Mean |  | 0.0755 | 16
| Median |  | 0.0755 | 0.0755
| Std Dev. |  | 0.00070711 | 30
| Rel.Std.Dev. |  | 0.9366%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 290 (Indicative Value 0.133 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.131 | 1
| 1 | 2 | 0.134 | 25
| Mean |  | 0.1325 | 16
| Median |  | 0.1325 | 0.1325
| Std Dev. |  | 0.0021213 | 31
| Rel.Std.Dev. |  | 1.601%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 290 (Indicative Value 0.925 wt.%) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.925 | 1
| 1 | 2 | 0.925 | 26
| Mean |  | 0.925 | 16
| Median |  | 0.925 | 0.925
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 290 (Indicative Value 67.94 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.92 | 1
| 1 | 2 | 67.95 | 23
| Mean |  | 67.935 | 16
| Median |  | 67.935 | 67.935
| Std Dev. |  | 0.021213 | 29
| Rel.Std.Dev. |  | 0.03123%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 290 (Indicative Value 1.31 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.313 | 24
| Mean |  | 1.3065 | 16
| Median |  | 1.3065 | 1.3065
| Std Dev. |  | 0.0091924 | 30
| Rel.Std.Dev. |  | 0.7036%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 290 (Indicative Value 0.73 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.73 | 1
| 1 | 2 | 0.73 | 25
| Mean |  | 0.73 | 16
| Median |  | 0.73 | 0.73
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 290 (Indicative Value 2.92 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.84 | 1
| 1 | 2 | 3 | 28
| Mean |  | 2.92 | 16
| Median |  | 2.92 | 2.92
| Std Dev. |  | 0.11314 | 34
| Rel.Std.Dev. |  | 3.875%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 290 (Indicative Value 0.215 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.21 | 1
| 1 | 2 | 0.22 | 30
| Mean |  | 0.215 | 16
| Median |  | 0.215 | 0.215
| Std Dev. |  | 0.0070711 | 36
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 290 (Indicative Value 0.51 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.52 | 30
| Mean |  | 0.51 | 16
| Median |  | 0.51 | 0.51
| Std Dev. |  | 0.014142 | 36
| Rel.Std.Dev. |  | 2.773%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 290 (Indicative Value 0.5 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.4 | 32
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 290 (Indicative Value 361 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 365 | 1
| 1 | 2 | 356 | 4
| Mean |  | 360.5 | 16
| Median |  | 360.5 | 360.5
| Std Dev. |  | 6.364 | 39
| Rel.Std.Dev. |  | 1.765%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 290 (Indicative Value 760 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 767 | 1
| 1 | 2 | 752 | 5
| Mean |  | 759.5 | 16
| Median |  | 759.5 | 759.5
| Std Dev. |  | 10.6066 | 40
| Rel.Std.Dev. |  | 1.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 290 (Indicative Value 2.3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.2 | 6
| Mean |  | 2.3 | 16
| Median |  | 2.3 | 2.3
| Std Dev. |  | 0.14142 | 41
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 290 (Indicative Value 0.43 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.44 | 1
| 1 | 2 | 0.42 | 7
| Mean |  | 0.43 | 16
| Median |  | 0.43 | 0.43
| Std Dev. |  | 0.014142 | 42
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 290 (Indicative Value < 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 9
| Mean |  |  | 16
| Median |  |  | < 0.1
| Std Dev. |  |  | 43
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 290 (Indicative Value 82 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 82.2 | 1
| 1 | 2 | 81.7 | 38
| Mean |  | 81.95 | 16
| Median |  | 81.95 | 81.95
| Std Dev. |  | 0.35355 | 44
| Rel.Std.Dev. |  | 0.4314%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 290 (Indicative Value 15.2 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.3 | 1
| 1 | 2 | 15 | 10
| Mean |  | 15.15 | 16
| Median |  | 15.15 | 15.15
| Std Dev. |  | 0.21213 | 45
| Rel.Std.Dev. |  | 1.4%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 290 (Indicative Value 134 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 132 | 1
| 1 | 2 | 136 | 11
| Mean |  | 134 | 16
| Median |  | 134 | 134
| Std Dev. |  | 2.8284 | 46
| Rel.Std.Dev. |  | 2.111%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 290 (Indicative Value 11.2 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.7 | 1
| 1 | 2 | 10.6 | 41
| Mean |  | 11.15 | 16
| Median |  | 11.15 | 11.15
| Std Dev. |  | 0.77782 | 47
| Rel.Std.Dev. |  | 6.976%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 290 (Indicative Value 40 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 40 | 12
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 0 | 48
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 290 (Indicative Value 5.95 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.95 | 1
| 1 | 2 | 5.94 | 18
| Mean |  | 5.945 | 16
| Median |  | 5.945 | 5.945
| Std Dev. |  | 0.0070711 | 49
| Rel.Std.Dev. |  | 0.1189%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 290 (Indicative Value 3.59 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.74 | 1
| 1 | 2 | 3.43 | 19
| Mean |  | 3.585 | 16
| Median |  | 3.585 | 3.585
| Std Dev. |  | 0.2192 | 50
| Rel.Std.Dev. |  | 6.114%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 290 (Indicative Value 1.39 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.44 | 1
| 1 | 2 | 1.34 | 20
| Mean |  | 1.39 | 16
| Median |  | 1.39 | 1.39
| Std Dev. |  | 0.070711 | 51
| Rel.Std.Dev. |  | 5.087%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 290 (Indicative Value 19.4 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.3 | 1
| 1 | 2 | 18.5 | 46
| Mean |  | 19.4 | 16
| Median |  | 19.4 | 19.4
| Std Dev. |  | 1.2728 | 52
| Rel.Std.Dev. |  | 6.561%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 290 (Indicative Value 6.07 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.36 | 1
| 1 | 2 | 5.78 | 21
| Mean |  | 6.07 | 16
| Median |  | 6.07 | 6.07
| Std Dev. |  | 0.41012 | 53
| Rel.Std.Dev. |  | 6.757%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 290 (Indicative Value 1.48 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.45 | 1
| 1 | 2 | 1.5 | 48
| Mean |  | 1.475 | 16
| Median |  | 1.475 | 1.475
| Std Dev. |  | 0.035355 | 54
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 290 (Indicative Value 7.04 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.07 | 1
| 1 | 2 | 7.01 | 32
| Mean |  | 7.04 | 16
| Median |  | 7.04 | 7.04
| Std Dev. |  | 0.042426 | 38
| Rel.Std.Dev. |  | 0.6026%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 290 (Indicative Value 1.25 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.25 | 17
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 290 (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 34
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 290 (Indicative Value 42 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 42.7 | 1
| 1 | 2 | 41.3 | 5
| Mean |  | 42 | 16
| Median |  | 42 | 42
| Std Dev. |  | 0.98995 | 41
| Rel.Std.Dev. |  | 2.357%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 290 (Indicative Value 0.45 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.43 | 1
| 1 | 2 | 0.47 | 18
| Mean |  | 0.45 | 16
| Median |  | 0.45 | 0.45
| Std Dev. |  | 0.028284 | 42
| Rel.Std.Dev. |  | 6.285%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 290 (Indicative Value 0.059 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0601 | 1
| 1 | 2 | 0.0573 | 8
| Mean |  | 0.0587 | 16
| Median |  | 0.0587 | 0.0587
| Std Dev. |  | 0.0019799 | 43
| Rel.Std.Dev. |  | 3.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 290 (Indicative Value 0.9 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 1 | 9
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 290 (Indicative Value 15.2 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.4 | 1
| 1 | 2 | 15 | 10
| Mean |  | 15.2 | 16
| Median |  | 15.2 | 15.2
| Std Dev. |  | 0.28284 | 45
| Rel.Std.Dev. |  | 1.861%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 290 (Indicative Value 37.1 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 37.8 | 1
| 1 | 2 | 36.4 | 19
| Mean |  | 37.1 | 16
| Median |  | 37.1 | 37.1
| Std Dev. |  | 0.98995 | 46
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 290 (Indicative Value 80 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84 | 1
| 1 | 2 | 76 | 11
| Mean |  | 80 | 16
| Median |  | 80 | 80
| Std Dev. |  | 5.6569 | 47
| Rel.Std.Dev. |  | 7.071%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 290 (Indicative Value 30.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31 | 1
| 1 | 2 | 30 | 13
| Mean |  | 30.5 | 16
| Median |  | 30.5 | 30.5
| Std Dev. |  | 0.70711 | 48
| Rel.Std.Dev. |  | 2.318%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 290 (Indicative Value 9.71 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.9 | 1
| 1 | 2 | 9.52 | 21
| Mean |  | 9.71 | 16
| Median |  | 9.71 | 9.71
| Std Dev. |  | 0.2687 | 49
| Rel.Std.Dev. |  | 2.767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 290 (Indicative Value 172 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 177 | 1
| 1 | 2 | 166 | 44
| Mean |  | 171.5 | 16
| Median |  | 171.5 | 171.5
| Std Dev. |  | 7.7782 | 50
| Rel.Std.Dev. |  | 4.535%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 290 (Indicative Value 0.04 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.06 | 45
| Mean |  | 0.04 | 16
| Median |  | 0.04 | 0.04
| Std Dev. |  | 0.028284 | 51
| Rel.Std.Dev. |  | 70.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 290 (Indicative Value 0.939 wt.%) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.948 | 1
| 1 | 2 | 0.93 | 26
| Mean |  | 0.939 | 16
| Median |  | 0.939 | 0.939
| Std Dev. |  | 0.012728 | 52
| Rel.Std.Dev. |  | 1.355%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 290 (Indicative Value 14.4 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.1 | 1
| 1 | 2 | 14.7 | 47
| Mean |  | 14.4 | 16
| Median |  | 14.4 | 14.4
| Std Dev. |  | 0.42426 | 53
| Rel.Std.Dev. |  | 2.946%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 290 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 48
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 54
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 290 (Indicative Value 7.59 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.59 | 1
| 1 | 2 | 7.59 | 18
| Mean |  | 7.59 | 16
| Median |  | 7.59 | 7.59
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 290 (Indicative Value 4.4 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.4 | 6
| Mean |  | 4.4 | 16
| Median |  | 4.4 | 4.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 290 (Indicative Value 101 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 103 | 1
| 1 | 2 | 98.6 | 7
| Mean |  | 100.8 | 16
| Median |  | 100.8 | 100.8
| Std Dev. |  | 3.1113 | 40
| Rel.Std.Dev. |  | 3.087%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 290 (Indicative Value 1.18 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.15 | 35
| Mean |  | 1.175 | 16
| Median |  | 1.175 | 1.175
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 290 (Indicative Value 1.01 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.01 | 1
| 1 | 2 | 1 | 19
| Mean |  | 1.005 | 16
| Median |  | 1.005 | 1.005
| Std Dev. |  | 0.0070711 | 42
| Rel.Std.Dev. |  | 0.7036%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 290 (Indicative Value 0.15 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | 0.2 | 37
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 43
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 290 (Indicative Value 16.2 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.4 | 1
| 1 | 2 | 16 | 38
| Mean |  | 16.2 | 16
| Median |  | 16.2 | 16.2
| Std Dev. |  | 0.28284 | 44
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 290 (Indicative Value 0.449 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.453 | 1
| 1 | 2 | 0.444 | 8
| Mean |  | 0.4485 | 16
| Median |  | 0.4485 | 0.4485
| Std Dev. |  | 0.006364 | 45
| Rel.Std.Dev. |  | 1.419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 290 (Indicative Value 1.1 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1 | 40
| Mean |  | 1.1 | 16
| Median |  | 1.1 | 1.1
| Std Dev. |  | 0.14142 | 46
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 290 (Indicative Value 0.55 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.51 | 20
| Mean |  | 0.55 | 16
| Median |  | 0.55 | 0.55
| Std Dev. |  | 0.056569 | 47
| Rel.Std.Dev. |  | 10.29%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 290 (Indicative Value 3.36 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.36 | 1
| 1 | 2 | 3.36 | 42
| Mean |  | 3.36 | 16
| Median |  | 3.36 | 3.36
| Std Dev. |  | 0 | 48
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 290 (Indicative Value 103 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 105 | 1
| 1 | 2 | 101 | 9
| Mean |  | 103 | 16
| Median |  | 103 | 103
| Std Dev. |  | 2.8284 | 49
| Rel.Std.Dev. |  | 2.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 290 (Indicative Value 3.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 3.5 | 10
| Mean |  | 3.75 | 16
| Median |  | 3.75 | 3.75
| Std Dev. |  | 0.35355 | 50
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 290 (Indicative Value 31.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32.3 | 1
| 1 | 2 | 30.9 | 11
| Mean |  | 31.6 | 16
| Median |  | 31.6 | 31.6
| Std Dev. |  | 0.98995 | 51
| Rel.Std.Dev. |  | 3.133%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 290 (Indicative Value 3.34 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.48 | 1
| 1 | 2 | 3.19 | 21
| Mean |  | 3.335 | 16
| Median |  | 3.335 | 3.335
| Std Dev. |  | 0.20506 | 52
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 290 (Indicative Value 93 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 85 | 12
| Mean |  | 92.5 | 16
| Median |  | 92.5 | 92.5
| Std Dev. |  | 10.6066 | 53
| Rel.Std.Dev. |  | 11.47%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 290 (Indicative Value 251 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 256 | 1
| 1 | 2 | 245 | 48
| Mean |  | 250.5 | 16
| Median |  | 250.5 | 250.5
| Std Dev. |  | 7.7782 | 54
| Rel.Std.Dev. |  | 3.105%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.