## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 555b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 12.3891 | 12.2632 | 12.5149 | 12.2672 | 12.5109
BaO, Barium oxide (ppm) | 316.1379 | 256.6397 | 375.6362 | 298.0396 | 334.2363
CaO, Calcium oxide (wt.%) | 0.20149 | 0.19372 | 0.20927 | 0.19288 | 0.21011
Co, Cobalt (wt.%) | 1.0848 | 1.0598 | 1.1098 | 1.0682 | 1.1015
Cu, Copper (wt.%) | 4.5728 | 4.4664 | 4.6792 | 4.5245 | 4.621
Fe2O3, Iron(III) oxide (wt.%) | 5.5509 | 5.4525 | 5.6492 | 5.4951 | 5.6067
K2O, Potassium oxide (wt.%) | 3.079 | 3.0449 | 3.1131 | 3.0508 | 3.1073
MgO, Magnesium oxide (wt.%) | 5.1339 | 5.0727 | 5.195 | 5.0842 | 5.1836
MnO, Manganese oxide (wt.%) | 0.11679 | 0.11026 | 0.12333 | 0.11368 | 0.11991
Ni, Nickel (ppm) | 232.5519 | 198.1168 | 266.987 | 213.59 | 251.5137
P2O5, Phosphorus(V) oxide (wt.%) | 0.083499 | 0.080055 | 0.086944 | 0.081515 | 0.085484
SiO2, Silicon dioxide (wt.%) | 51.7731 | 51.3129 | 52.2333 | 51.425 | 52.1212
SO3, Sulphur trioxide (wt.%) | 11.0536 | 10.5578 | 11.5493 | 10.9326 | 11.1746
Sr, Strontium (ppm) | 67.1921 | 47.1214 | 87.2628 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.78852 | 0.76535 | 0.81169 | 0.77336 | 0.80368
V2O5, Vanadium(V) oxide (ppm) | 691.7704 | 630.8738 | 752.6671 | 641.3511 | 742.1898
Zr, Zirconium (ppm) | 237.6548 | 226.5409 | 248.7687 | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 12.9193 | 12.7483 | 13.0902 | 12.8303 | 13.0083
Infrared Combustion
C, Carbon (wt.%) | 4.5302 | 4.432 | 4.6284 | 4.4883 | 4.5722
S, Sulphur (wt.%) | 4.4824 | 4.3971 | 4.5677 | 4.443 | 4.5218
Sulphuric Acid 5% Leach
Co, Cobalt (wt.%) | 0.43156 | 0.38837 | 0.47476 | 0.40957 | 0.45356
Cu, Copper (wt.%) | 1.1872 | 1.1462 | 1.2282 | 1.1677 | 1.2067
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.4231 | 6.2833 | 6.5629 | 6.3525 | 6.4937
As, Arsenic (ppm) | 65.7417 | 57.9889 | 73.4945 | 63.8937 | 67.5897
B, Boron (ppm) | 231.8282 | 209.3881 | 254.2682 | 219.7981 | 243.8583
Ba, Barium (ppm) | 258.7861 | 247.5186 | 270.0537 | 251.4609 | 266.1114
Be, Beryllium (ppm) | 5.3275 | 4.6412 | 6.0139 | IND | IND
Bi, Bismuth (ppm) | 1.5558 | 1.2795 | 1.8321 | IND | IND
Ca, Calcium (wt.%) | 0.14299 | 0.12364 | 0.16234 | 0.13016 | 0.15582
Ce, Cerium (ppm) | 160.8185 | 139.98 | 181.657 | 146.1036 | 175.5335
Co, Cobalt (wt.%) | 1.1053 | 1.0654 | 1.1453 | 1.0727 | 1.138
Cr, Chromium (ppm) | 111.7493 | 98.5631 | 124.9355 | 105.4504 | 118.0482
Cs, Caesium (ppm) | 2.259 | 1.9832 | 2.5349 | 2.0596 | 2.4584
Cu, Copper (wt.%) | 4.5665 | 4.4526 | 4.6804 | 4.5131 | 4.6199
Dy, Dysprosium (ppm) | 6.6976 | 6.1485 | 7.2467 | 6.2446 | 7.1506
Er, Erbium (ppm) | 3.8268 | 3.5217 | 4.1318 | 3.5888 | 4.0647
Eu, Europium (ppm) | 0.9668 | 0.85902 | 1.0746 | 0.91073 | 1.0229
Fe, Iron (wt.%) | 3.8742 | 3.7638 | 3.9847 | 3.8145 | 3.934
Ga, Gallium (ppm) | 29.1167 | 27.7925 | 30.441 | 27.6634 | 30.57
Gd, Gadolinium (ppm) | 7.257 | 6.3639 | 8.15 | 6.6958 | 7.8182
Ge, Germanium (ppm) | 2.0081 | 1.9546 | 2.0616 | IND | IND
Ho, Holmium (ppm) | 1.3779 | 1.2086 | 1.5471 | 1.28 | 1.4758
In, Indium (ppm) | 2.4674 | 2.1782 | 2.7565 | 2.3427 | 2.592
K, Potassium (wt.%) | 2.5809 | 2.4719 | 2.6899 | 2.5219 | 2.6399
La, Lanthanum (ppm) | 76.5578 | 71.3279 | 81.7878 | 72.2936 | 80.8221
Li, Lithium (ppm) | 59.7113 | 55.7843 | 63.6383 | 57.6723 | 61.7504
Lu, Lutetium (ppm) | 0.57565 | 0.43889 | 0.71241 | 0.51922 | 0.63208
Mg, Magnesium (wt.%) | 3.0941 | 2.9958 | 3.1923 | 3.0424 | 3.1457
Mn, Manganese (wt.%) | 0.0925 | 0.08752 | 0.09748 | 0.089853 | 0.095148
Mo, Molybdenum (ppm) | 17.2971 | 15.521 | 19.0733 | 16.5387 | 18.0556
Nb, Niobium (ppm) | 25.3883 | 24.002 | 26.7746 | 23.6507 | 27.126
Nd, Neodymium (ppm) | 67.539 | 62.749 | 72.3289 | 63.4983 | 71.5797
Ni, Nickel (ppm) | 228.1801 | 213.7315 | 242.6287 | 219.3647 | 236.9954
P, Phosphorus (wt.%) | 0.040115 | 0.029873 | 0.050356 | IND | IND
Pb, Lead (ppm) | < 100 | IND | IND | IND | IND
Pr, Praseodymium (ppm) | 18.6327 | 17.4958 | 19.7696 | 17.3849 | 19.8805
Rb, Rubidium (ppm) | 95.6344 | 90.6929 | 100.5759 | 92.695 | 98.5739
S, Sulphur (wt.%) | 4.4524 | 4.3341 | 4.5708 | 4.3745 | 4.5304
Sb, Antimony (ppm) | 4.0001 | 3.5868 | 4.4134 | IND | IND
Sc, Scandium (ppm) | 11.793 | 10.7851 | 12.8008 | IND | IND
Si, Silicon (wt.%) | 24.6187 | 24.1437 | 25.0938 | 24.3154 | 24.9221
Sm, Samarium (ppm) | 9.0676 | 8.1022 | 10.033 | 8.2409 | 9.8943
Sr, Strontium (ppm) | 74.9489 | 69.7767 | 80.1211 | 72.8964 | 77.0014
Ta, Tantalum (ppm) | 1.9644 | 1.7178 | 2.2109 | IND | IND
Tb, Terbium (ppm) | 1.1933 | 1.0709 | 1.3157 | 1.1092 | 1.2774
Th, Thorium (ppm) | 15.7059 | 15.0807 | 16.3311 | 15.089 | 16.3227
Ti, Titanium (wt.%) | 0.46429 | 0.45335 | 0.47523 | 0.45494 | 0.47364
Tl, Thallium (ppm) | 1.0205 | 0.91266 | 1.1283 | IND | IND
Tm, Thulium (ppm) | 0.56714 | 0.49877 | 0.6355 | 0.52381 | 0.61047
U, Uranium (ppm) | 9.5113 | 8.9946 | 10.0279 | 9.1585 | 9.864
V, Vanadium (ppm) | 378.7095 | 361.7838 | 395.6352 | 367.4149 | 390.0041
W, Tungsten (ppm) | 2.9367 | 1.5979 | 4.2756 | IND | IND
Y, Yttrium (ppm) | 36.5396 | 34.583 | 38.4963 | 35.0966 | 37.9826
Yb, Ytterbium (ppm) | 3.4947 | 3.0735 | 3.9159 | 3.2527 | 3.7367
Zn, Zinc (ppm) | 31.6644 | 24.146 | 39.1828 | 29.262 | 34.0668
4-Acid Digestion
Ag, Silver (ppm) | 0.097963 | 0.073357 | 0.12257 | IND | IND
Al, Aluminium (wt.%) | 6.4043 | 6.2873 | 6.5212 | 6.2747 | 6.5338
As, Arsenic (ppm) | 63.9806 | 60.7321 | 67.2291 | 61.0503 | 66.911
Ba, Barium (ppm) | 254.294 | 244.5083 | 264.0797 | 248.3275 | 260.2605
Be, Beryllium (ppm) | 4.7558 | 4.6096 | 4.902 | 4.5965 | 4.9151
Bi, Bismuth (ppm) | 1.4754 | 1.413 | 1.5379 | 1.4353 | 1.5156
Ca, Calcium (wt.%) | 0.14453 | 0.13812 | 0.15093 | 0.14041 | 0.14864
Cd, Cadmium (ppm) | 0.45646 | 0.41911 | 0.4938 | 0.42853 | 0.48438
Ce, Cerium (ppm) | 148.4917 | 142.4809 | 154.5025 | 143.9353 | 153.0481
Co, Cobalt (wt.%) | 1.102 | 1.0723 | 1.1317 | 1.0783 | 1.1257
Cr, Chromium (ppm) | 93.2783 | 89.3996 | 97.157 | 89.8542 | 96.7024
Cs, Caesium (ppm) | 2.0321 | 1.9422 | 2.1219 | 1.9654 | 2.0987
Cu, Copper (wt.%) | 4.5807 | 4.4674 | 4.694 | 4.4956 | 4.6658
Dy, Dysprosium (ppm) | 4.6792 | 4.0172 | 5.3412 | 4.4071 | 4.9512
Er, Erbium (ppm) | 2.7576 | 2.3933 | 3.1219 | 2.5982 | 2.917
Eu, Europium (ppm) | 0.91493 | 0.85697 | 0.9729 | 0.86185 | 0.96802
Fe, Iron (wt.%) | 3.8319 | 3.7252 | 3.9386 | 3.7598 | 3.9041
Ga, Gallium (ppm) | 28.3601 | 27.2476 | 29.4727 | 27.6566 | 29.0637
Gd, Gadolinium (ppm) | 5.1774 | 4.4202 | 5.9347 | 4.9126 | 5.4423
Ge, Germanium (ppm) | 0.19654 | 0.15739 | 0.23569 | 0.1742 | 0.21888
Hf, Hafnium (ppm) | 5.3705 | 5.1384 | 5.6025 | 5.186 | 5.555
Ho, Holmium (ppm) | 0.902 | 0.7928 | 1.0112 | 0.85328 | 0.95071
In, Indium (ppm) | 2.506 | 2.3522 | 2.6598 | 2.425 | 2.5871
K, Potassium (wt.%) | 2.5371 | 2.4762 | 2.598 | 2.4758 | 2.5983
La, Lanthanum (ppm) | 71.7949 | 68.228 | 75.3619 | 69.6966 | 73.8933
Li, Lithium (ppm) | 59.3934 | 57.6373 | 61.1495 | 58.017 | 60.7697
Lu, Lutetium (ppm) | 0.44151 | 0.40575 | 0.47727 | 0.41809 | 0.46493
Mg, Magnesium (wt.%) | 3.0723 | 2.9918 | 3.1529 | 3.0037 | 3.1409
Mn, Manganese (wt.%) | 0.089023 | 0.085421 | 0.092626 | 0.087169 | 0.090878
Mo, Molybdenum (ppm) | 17.0696 | 16.3698 | 17.7694 | 16.5625 | 17.5767
Na, Sodium (wt.%) | 0.051973 | 0.048517 | 0.055429 | 0.049158 | 0.054788
Nb, Niobium (ppm) | 17.0885 | 14.9019 | 19.2752 | 16.5454 | 17.6316
Nd, Neodymium (ppm) | 64.0192 | 61.5753 | 66.4631 | 61.9587 | 66.0796
Ni, Nickel (ppm) | 220.1353 | 211.5211 | 228.7494 | 213.9411 | 226.3295
P, Phosphorus (wt.%) | 0.034577 | 0.032874 | 0.03628 | 0.033366 | 0.035789
Pb, Lead (ppm) | 13.6067 | 12.7662 | 14.4472 | 13.1837 | 14.0298
Pr, Praseodymium (ppm) | 16.9339 | 15.85 | 18.0178 | 16.5115 | 17.3563
Rb, Rubidium (ppm) | 94.4528 | 88.8597 | 100.0459 | 90.725 | 98.1806
Re, Rhenium (ppm) | 0.029874 | 0.026524 | 0.033225 | 0.026671 | 0.033077
S, Sulphur (wt.%) | 4.3937 | 4.2875 | 4.5 | 4.3127 | 4.4747
Sb, Antimony (ppm) | 3.7211 | 3.5001 | 3.9421 | 3.5614 | 3.8807
Sc, Scandium (ppm) | 12.5862 | 12.0055 | 13.167 | 12.2154 | 12.957
Se, Selenium (ppm) | 6.9708 | 5.9918 | 7.9498 | 6.5546 | 7.387
Sm, Samarium (ppm) | 8.8742 | 8.2481 | 9.5003 | 8.5061 | 9.2424
Sn, Tin (ppm) | 2.9872 | 2.7855 | 3.1889 | 2.847 | 3.1273
Sr, Strontium (ppm) | 69.8607 | 67.4603 | 72.2611 | 68.3937 | 71.3277
Ta, Tantalum (ppm) | 1.192 | 1.0543 | 1.3297 | 1.1373 | 1.2468
Tb, Terbium (ppm) | 0.81401 | 0.64505 | 0.98297 | 0.77141 | 0.85661
Te, Tellurium (ppm) | 0.059818 | 0.023703 | 0.095932 | IND | IND
Th, Thorium (ppm) | 15.977 | 15.2652 | 16.6888 | 15.4121 | 16.5419
Ti, Titanium (wt.%) | 0.27375 | 0.23699 | 0.31052 | 0.26515 | 0.28235
Tl, Thallium (ppm) | 0.9867 | 0.94181 | 1.0316 | 0.95508 | 1.0183
Tm, Thulium (ppm) | 0.41895 | 0.35643 | 0.48148 | 0.39526 | 0.44265
U, Uranium (ppm) | 9.2368 | 8.6937 | 9.7799 | 8.9881 | 9.4855
V, Vanadium (ppm) | 356.6079 | 344.4172 | 368.7987 | 348.648 | 364.5678
W, Tungsten (ppm) | 2.7793 | 2.5958 | 2.9629 | 2.6757 | 2.8829
Y, Yttrium (ppm) | 24.4386 | 23.4049 | 25.4723 | 23.7365 | 25.1407
Yb, Ytterbium (ppm) | 2.8839 | 2.6098 | 3.158 | 2.7603 | 3.0074
Zn, Zinc (ppm) | 29.0115 | 27.5071 | 30.5159 | 27.7459 | 30.2771
Zr, Zirconium (ppm) | 187.9751 | 182.7068 | 193.2434 | 183.1434 | 192.8068
Aqua Regia Digestion
Ag, Silver (ppm) | 0.075622 | 0.060483 | 0.090761 | IND | IND
Al, Aluminium (wt.%) | 0.9585 | 0.86767 | 1.0493 | 0.92247 | 0.99454
As, Arsenic (ppm) | 61.6015 | 59.2601 | 63.9429 | 60.0976 | 63.1054
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 56.6609 | 51.4195 | 61.9023 | 54.6409 | 58.681
Be, Beryllium (ppm) | 2.2126 | 2.052 | 2.3732 | 2.1441 | 2.2811
Bi, Bismuth (ppm) | 1.3874 | 1.3254 | 1.4494 | 1.3506 | 1.4241
Ca, Calcium (wt.%) | 0.13664 | 0.12929 | 0.14399 | 0.13363 | 0.13965
Cd, Cadmium (ppm) | 0.46427 | 0.41943 | 0.50912 | 0.43618 | 0.49237
Ce, Cerium (ppm) | 43.6838 | 39.5317 | 47.8359 | 42.0109 | 45.3568
Co, Cobalt (wt.%) | 1.0756 | 1.0426 | 1.1086 | 1.0533 | 1.098
Cr, Chromium (ppm) | 35.2455 | 33.3058 | 37.1852 | 34.1532 | 36.3379
Cs, Caesium (ppm) | 0.32386 | 0.26596 | 0.38176 | 0.30528 | 0.34244
Cu, Copper (wt.%) | 4.5716 | 4.5046 | 4.6385 | 4.515 | 4.6281
Dy, Dysprosium (ppm) | 0.99577 | 0.84939 | 1.1422 | 0.95869 | 1.0328
Er, Erbium (ppm) | 0.4069 | 0.2912 | 0.52259 | 0.38431 | 0.42948
Eu, Europium (ppm) | 0.38641 | 0.33639 | 0.43644 | 0.35906 | 0.41376
Fe, Iron (wt.%) | 3.6635 | 3.5565 | 3.7704 | 3.5806 | 3.7463
Ga, Gallium (ppm) | 8.2218 | 7.4463 | 8.9973 | 7.9223 | 8.5213
Gd, Gadolinium (ppm) | 2.025 | 1.8205 | 2.2295 | 1.902 | 2.1481
Ge, Germanium (ppm) | 0.095476 | 0.066005 | 0.12495 | IND | IND
Hf, Hafnium (ppm) | 0.33039 | 0.30015 | 0.36063 | 0.3115 | 0.34927
Hg, Mercury (ppm) | 0.060264 | 0.044461 | 0.076067 | IND | IND
Ho, Holmium (ppm) | 0.16772 | 0.12888 | 0.20655 | IND | IND
In, Indium (ppm) | 1.827 | 1.7504 | 1.9035 | 1.7909 | 1.863
K, Potassium (wt.%) | 0.19566 | 0.17975 | 0.21158 | 0.18167 | 0.20966
La, Lanthanum (ppm) | 20.3761 | 18.8566 | 21.8957 | 19.4174 | 21.3349
Li, Lithium (ppm) | 35.3459 | 31.7054 | 38.9864 | 34.0223 | 36.6695
Mg, Magnesium (wt.%) | 2.4173 | 2.3007 | 2.5338 | 2.3558 | 2.4787
Mn, Manganese (wt.%) | 0.088654 | 0.085267 | 0.092041 | 0.085881 | 0.091427
Mo, Molybdenum (ppm) | 16.2231 | 15.6568 | 16.7893 | 15.8809 | 16.5652
Na, Sodium (wt.%) | 0.010713 | 0.0093532 | 0.012074 | 0.0099922 | 0.011435
Nb, Niobium (ppm) | 0.093611 | 0.07362 | 0.1136 | IND | IND
Nd, Neodymium (ppm) | 22.9844 | 18.739 | 27.2298 | 21.3555 | 24.6133
Ni, Nickel (ppm) | 213.8519 | 206.3734 | 221.3303 | 208.4502 | 219.2535
P, Phosphorus (wt.%) | 0.020017 | 0.018817 | 0.021218 | 0.018936 | 0.021098
Pb, Lead (ppm) | 10.9141 | 10.3176 | 11.5106 | 10.5841 | 11.244
Pr, Praseodymium (ppm) | 5.4327 | 3.8134 | 7.0519 | 5.2281 | 5.6373
Rb, Rubidium (ppm) | 6.4953 | 5.7628 | 7.2279 | 6.0874 | 6.9033
Re, Rhenium (ppm) | 0.026561 | 0.024021 | 0.029101 | 0.024654 | 0.028468
S, Sulphur (wt.%) | 4.3112 | 4.1919 | 4.4306 | 4.2209 | 4.4016
Sb, Antimony (ppm) | 2.1662 | 1.9194 | 2.413 | 2.0527 | 2.2797
Sc, Scandium (ppm) | 3.2952 | 3.0242 | 3.5662 | 3.1298 | 3.4606
Se, Selenium (ppm) | 6.6992 | 6.0816 | 7.3167 | 6.48 | 6.9183
Sm, Samarium (ppm) | 3.5079 | 2.7952 | 4.2205 | 3.31 | 3.7058
Sn, Tin (ppm) | 1.5356 | 1.4176 | 1.6536 | 1.4738 | 1.5974
Sr, Strontium (ppm) | 15.1067 | 13.6911 | 16.5222 | 14.3971 | 15.8162
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.20531 | 0.15641 | 0.2542 | 0.19451 | 0.21611
Te, Tellurium (ppm) | 0.053516 | 0.032717 | 0.074314 | IND | IND
Th, Thorium (ppm) | 6.3228 | 5.7736 | 6.872 | 6.0029 | 6.6426
U, Uranium (ppm) | 3.491 | 3.3146 | 3.6674 | 3.3626 | 3.6194
V, Vanadium (ppm) | 62.0724 | 57.5403 | 66.6045 | 59.7 | 64.4448
W, Tungsten (ppm) | 1.0849 | 0.99067 | 1.1791 | 1.0204 | 1.1494
Y, Yttrium (ppm) | 3.9366 | 3.6684 | 4.2048 | 3.786 | 4.0872
Yb, Ytterbium (ppm) | 0.31535 | 0.21327 | 0.41742 | 0.29562 | 0.33507
Zn, Zinc (ppm) | 26.6028 | 25.3336 | 27.8721 | 25.646 | 27.5596
Zr, Zirconium (ppm) | 10.1145 | 9.3496 | 10.8794 | 9.7133 | 10.5157
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 555b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Ag | ppm | 0.0345 | HfO2 | ppm | 252.9417 | Sc | ppm | 13.8577
As | ppm | 60.2222 | Hg | ppm | < 100 | Se | ppm | 2.6667
Be | ppm | 3 | Ho | ppm | 1.387 | Sm | ppm | 9.2526
Bi | ppm | < 100 | In | ppm | 2.7833 | Sn | ppm | 93.2292
Cd | ppm | < 100 | La | ppm | 77.6183 | Ta | ppm | 1.4667
Ce | ppm | 129.9572 | Lu | ppm | 0.57725 | Tb | ppm | 1.1494
Cl | ppm | 223.8765 | Mo | ppm | < 50 | Te | ppm | 10.9167
Cr2O3 | ppm | 159.3327 | Na2O | wt.% | 0.10333 | Th | ppm | 15.6751
Cs | ppm | 2.2167 | Nb | ppm | 40.0056 | Tl | ppm | 0.79667
Dy | ppm | 6.7094 | Nd | ppm | 62.5805 | Tm | ppm | 0.56608
Er | ppm | 3.8665 | Pb | ppm | < 100 | U | ppm | 9.1254
Eu | ppm | 0.89117 | Pr | ppm | 17.3803 | W | ppm | 12.66
Ga | ppm | 28.5833 | Rb | ppm | 103.0667 | Y | ppm | 33.9574
Gd | ppm | 7.1144 | Re | ppm | 18.2833 | Yb | ppm | 3.6357
Ge | ppm | < 100 | Sb | ppm | < 50 | Zn | ppm | 43.6111
Peroxide Fusion ICP
Ag | ppm | 0.42583 | Na | wt.% | 0.05 | Te | ppm | < 2
Cd | ppm | < 10 | Re | ppm | 0.059167 | Zr | ppm | 246.5314
Hf | ppm | 6.8398 | Se | ppm | 10.5
Hg | ppm | < 0.1 | Sn | ppm | 3.4735
4-Acid Digestion
Hg | ppm | 0.073333 | Pt | ppb | 10.5
Aqua Regia Digestion
Ir | ppm | < 0.003 | Pt | ppb | < 5 | Tm | ppm | 0.050722
Lu | ppm | 0.057506 | Ti | wt.% | 0.0015157
Pd | ppb | < 10 | Tl | ppm | 0.35805
3-Acid Digestion (no HF)
Ag | ppm | 0.38333 | Gd | ppm | 5.6333 | S | wt.% | 4.1867
Al2O3 | wt.% | 12.1667 | Hf | ppm | 5.9667 | Sc | ppm | 12.95
Ba | ppm | 247.1667 | Ho | ppm | 0.92667 | Sm | ppm | 8.1167
Be | ppm | 5.85 | K2O | wt.% | 3.095 | Sn | ppm | 2.97
Bi | ppm | 1.95 | La | ppm | 78.6667 | Sr | ppm | 71.5
CaO | wt.% | 0.19 | Li | ppm | 61.5 | Ta | ppm | 0.35333
Cd | ppm | 0.58333 | MgO | wt.% | 5.12 | Tb | ppm | 0.73667
Ce | ppm | 142 | MnO | wt.% | 0.1145 | Th | ppm | 17.6
Co | wt.% | 1.17 | Mo | ppm | 19.5667 | TiO2 | wt.% | 0.78233
Cr | ppm | 104.1667 | Na2O | wt.% | 0.073333 | U | ppm | 10.4867
Cs | ppm | 2.15 | Nb | ppm | 23.35 | V | ppm | 374.1667
Cu | wt.% | 4.8383 | Nd | ppm | 64.5667 | W | ppm | 3.7
Dy | ppm | 4.8267 | Ni | ppm | 253.6667 | Y | ppm | 27.7333
Er | ppm | 2.8983 | P2O5 | wt.% | < 0.002 | Yb | ppm | 2.8967
Eu | ppm | 0.95833 | Pb | ppm | 15.8167 | Zn | ppm | 66.8333
Fe2O3 | wt.% | 5.4083 | Pr | ppm | 17.3833 | Zr | ppm | 197.6667
Ga | ppm | 28.3 | Rb | ppm | 102.5167
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 555b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 12.3891 | 0.18351 | 12.0221 | 12.7561 | 11.8385 | 12.9396 | 1.481% | 2.962% | 4.444% | 11.7696 | 13.0085
BaO, ppm | 316.1379 | 62.7136 | 190.7107 | 441.5652 | 127.997 | 504.2789 | 19.84% | 39.67% | 59.51% | 300.331 | 331.9448
CaO, wt.% | 0.20149 | 0.0065515 | 0.18839 | 0.2146 | 0.18184 | 0.22115 | 3.251% | 6.503% | 9.754% | 0.19142 | 0.21157
Co, wt.% | 1.0848 | 0.03605 | 1.0127 | 1.1569 | 0.97668 | 1.193 | 3.323% | 6.646% | 9.969% | 1.0306 | 1.1391
Cu, wt.% | 4.5728 | 0.18319 | 4.2064 | 4.9391 | 4.0232 | 5.1223 | 4.006% | 8.012% | 12.02% | 4.3441 | 4.8014
Fe2O3, wt.% | 5.5509 | 0.17632 | 5.1982 | 5.9035 | 5.0219 | 6.0798 | 3.176% | 6.353% | 9.529% | 5.2733 | 5.8284
K2O, wt.% | 3.079 | 0.052153 | 2.9747 | 3.1833 | 2.9225 | 3.2355 | 1.694% | 3.388% | 5.082% | 2.9251 | 3.233
MgO, wt.% | 5.1339 | 0.11077 | 4.9123 | 5.3554 | 4.8016 | 5.4662 | 2.158% | 4.315% | 6.473% | 4.8772 | 5.3906
MnO, wt.% | 0.11679 | 0.0045901 | 0.10761 | 0.12597 | 0.10302 | 0.13056 | 3.93% | 7.86% | 11.79% | 0.11095 | 0.12263
Ni, ppm | 232.5519 | 27.0521 | 178.4477 | 286.6561 | 151.3956 | 313.7081 | 11.63% | 23.27% | 34.9% | 220.9243 | 244.1795
P2O5, wt.% | 0.083499 | 0.0062603 | 0.070979 | 0.09602 | 0.064719 | 0.10228 | 7.497% | 14.99% | 22.49% | 0.079324 | 0.087674
SiO2, wt.% | 51.7731 | 0.67563 | 50.4219 | 53.1244 | 49.7462 | 53.8 | 1.305% | 2.61% | 3.915% | 49.1844 | 54.3618
SO3, wt.% | 11.0536 | 0.59092 | 9.8717 | 12.2354 | 9.2808 | 12.8264 | 5.346% | 10.69% | 16.04% | 10.5009 | 11.6063
Sr, ppm | 67.1921 | 19.9078 | 27.3765 | 107.0078 | 7.4686 | 126.9156 | 29.63% | 59.26% | 88.88% | 63.8325 | 70.5517
TiO2, wt.% | 0.78852 | 0.036318 | 0.71588 | 0.86115 | 0.67956 | 0.89747 | 4.606% | 9.212% | 13.82% | 0.74909 | 0.82794
V2O5, ppm | 691.7704 | 62.1271 | 567.5162 | 816.0246 | 505.3891 | 878.1518 | 8.981% | 17.96% | 26.94% | 657.1819 | 726.359
Zr, ppm | 237.6548 | 27.2239 | 183.2071 | 292.1025 | 155.9832 | 319.3264 | 11.46% | 22.91% | 34.37% | 225.7721 | 249.5375
Thermogravimetry
LOI1000, wt.% | 12.9193 | 0.34381 | 12.2317 | 13.6069 | 11.8878 | 13.9507 | 2.661% | 5.322% | 7.984% | 12.2733 | 13.5652
Infrared Combustion
C, wt.% | 4.5302 | 0.18535 | 4.1596 | 4.9009 | 3.9742 | 5.0863 | 4.091% | 8.183% | 12.27% | 4.3037 | 4.7568
S, wt.% | 4.4824 | 0.095574 | 4.2913 | 4.6736 | 4.1957 | 4.7691 | 2.132% | 4.264% | 6.397% | 4.2583 | 4.7065
Sulphuric Acid 5% Leach
Co, wt.% | 0.43156 | 0.035898 | 0.35977 | 0.50336 | 0.32387 | 0.53926 | 8.318% | 16.64% | 24.95% | 0.40999 | 0.45314
Cu, wt.% | 1.1872 | 0.051971 | 1.0833 | 1.2911 | 1.0313 | 1.3431 | 4.378% | 8.755% | 13.13% | 1.1278 | 1.2466
Peroxide Fusion ICP
Al, wt.% | 6.4231 | 0.17986 | 6.0634 | 6.7828 | 5.8835 | 6.9627 | 2.8% | 5.6% | 8.401% | 6.1019 | 6.7442
As, ppm | 65.7417 | 5.7905 | 54.1608 | 77.3226 | 48.3703 | 83.1131 | 8.808% | 17.62% | 26.42% | 62.4546 | 69.0288
B, ppm | 231.8282 | 29.3134 | 173.2013 | 290.455 | 143.8878 | 319.7685 | 12.64% | 25.29% | 37.93% | 220.2368 | 243.4196
Ba, ppm | 258.7861 | 12.4724 | 233.8413 | 283.7309 | 221.369 | 296.2033 | 4.82% | 9.639% | 14.46% | 245.8468 | 271.7255
Be, ppm | 5.3275 | 0.38778 | 4.552 | 6.1031 | 4.1642 | 6.4909 | 7.279% | 14.56% | 21.84% | 5.0612 | 5.5939
Bi, ppm | 1.5558 | 0.22567 | 1.1044 | 2.0071 | 0.87875 | 2.2328 | 14.51% | 29.01% | 43.52% | 1.478 | 1.6336
Ca, wt.% | 0.14299 | 0.026177 | 0.090637 | 0.19535 | 0.064459 | 0.22152 | 18.31% | 36.61% | 54.92% | 0.13584 | 0.15014
Ce, ppm | 160.8185 | 15.539 | 129.7405 | 191.8965 | 114.2015 | 207.4355 | 9.662% | 19.32% | 28.99% | 152.7776 | 168.8594
Co, wt.% | 1.1053 | 0.034432 | 1.0365 | 1.1742 | 1.002 | 1.2086 | 3.115% | 6.23% | 9.345% | 1.0501 | 1.1606
Cr, ppm | 111.7493 | 22.7606 | 66.2282 | 157.2704 | 43.4676 | 180.031 | 20.37% | 40.74% | 61.1% | 106.1618 | 117.3367
Cs, ppm | 2.259 | 0.21465 | 1.8297 | 2.6883 | 1.6151 | 2.903 | 9.502% | 19% | 28.51% | 2.1461 | 2.372
Cu, wt.% | 4.5665 | 0.14091 | 4.2847 | 4.8483 | 4.1438 | 4.9893 | 3.086% | 6.172% | 9.257% | 4.3382 | 4.7948
Dy, ppm | 6.6976 | 0.45762 | 5.7824 | 7.6128 | 5.3247 | 8.0705 | 6.833% | 13.67% | 20.5% | 6.3627 | 7.0325
Er, ppm | 3.8268 | 0.22131 | 3.3841 | 4.2694 | 3.1628 | 4.4907 | 5.783% | 11.57% | 17.35% | 3.6354 | 4.0181
Eu, ppm | 0.9668 | 0.070061 | 0.82668 | 1.1069 | 0.75662 | 1.177 | 7.247% | 14.49% | 21.74% | 0.91846 | 1.0151
Fe, wt.% | 3.8742 | 0.14323 | 3.5878 | 4.1607 | 3.4446 | 4.3039 | 3.697% | 7.394% | 11.09% | 3.6805 | 4.068
Ga, ppm | 29.1167 | 1.4164 | 26.284 | 31.9495 | 24.8676 | 33.3658 | 4.864% | 9.729% | 14.59% | 27.6609 | 30.5726
Gd, ppm | 7.257 | 0.72588 | 5.8052 | 8.7087 | 5.0794 | 9.4346 | 10% | 20% | 30.01% | 6.8941 | 7.6198
Ge, ppm | 2.0081 | 0.049221 | 1.9097 | 2.1066 | 1.8605 | 2.1558 | 2.451% | 4.902% | 7.353% | 1.9077 | 2.1085
Ho, ppm | 1.3779 | 0.099772 | 1.1784 | 1.5774 | 1.0786 | 1.6772 | 7.241% | 14.48% | 21.72% | 1.309 | 1.4468
In, ppm | 2.4674 | 0.14928 | 2.1688 | 2.7659 | 2.0195 | 2.9152 | 6.05% | 12.1% | 18.15% | 2.344 | 2.5907
K, wt.% | 2.5809 | 0.14488 | 2.2911 | 2.8706 | 2.1462 | 3.0155 | 5.614% | 11.23% | 16.84% | 2.4518 | 2.7099
La, ppm | 76.5578 | 6.3996 | 63.7586 | 89.3571 | 57.3589 | 95.7567 | 8.359% | 16.72% | 25.08% | 72.7299 | 80.3857
Li, ppm | 59.7113 | 4.5442 | 50.623 | 68.7997 | 46.0789 | 73.3438 | 7.61% | 15.22% | 22.83% | 56.7258 | 62.6969
Lu, ppm | 0.57565 | 0.09301 | 0.38963 | 0.76167 | 0.29662 | 0.85468 | 16.16% | 32.31% | 48.47% | 0.54687 | 0.60443
Mg, wt.% | 3.0941 | 0.11235 | 2.8694 | 3.3187 | 2.757 | 3.4311 | 3.631% | 7.262% | 10.89% | 2.9394 | 3.2488
Mn, wt.% | 0.0925 | 0.0037378 | 0.085024 | 0.099976 | 0.081287 | 0.10371 | 4.041% | 8.082% | 12.12% | 0.087875 | 0.097125
Mo, ppm | 17.2971 | 1.9634 | 13.3704 | 21.2239 | 11.407 | 23.1873 | 11.35% | 22.7% | 34.05% | 16.4323 | 18.162
Nb, ppm | 25.3883 | 1.0103 | 23.3677 | 27.4089 | 22.3574 | 28.4193 | 3.979% | 7.959% | 11.94% | 24.1189 | 26.6577
Nd, ppm | 67.539 | 4.3271 | 58.8847 | 76.1932 | 54.5576 | 80.5203 | 6.407% | 12.81% | 19.22% | 64.162 | 70.9159
Ni, ppm | 228.1801 | 21.695 | 184.7901 | 271.57 | 163.0951 | 293.265 | 9.508% | 19.02% | 28.52% | 216.7711 | 239.5891
P, wt.% | 0.040115 | 0.0085088 | 0.023097 | 0.057132 | 0.014588 | 0.065641 | 21.21% | 42.42% | 63.63% | 0.038109 | 0.04212
Pb, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Pr, ppm | 18.6327 | 0.89123 | 16.8502 | 20.4151 | 15.959 | 21.3064 | 4.783% | 9.566% | 14.35% | 17.701 | 19.5643
Rb, ppm | 95.6344 | 5.148 | 85.3384 | 105.9304 | 80.1904 | 111.0784 | 5.383% | 10.77% | 16.15% | 90.8527 | 100.4162
S, wt.% | 4.4524 | 0.14378 | 4.1649 | 4.74 | 4.0211 | 4.8838 | 3.229% | 6.458% | 9.688% | 4.2298 | 4.6751
Sb, ppm | 4.0001 | 0.13988 | 3.7203 | 4.2798 | 3.5804 | 4.4197 | 3.497% | 6.994% | 10.49% | 3.8001 | 4.2001
Sc, ppm | 11.793 | 1.0101 | 9.7728 | 13.8132 | 8.7627 | 14.8232 | 8.565% | 17.13% | 25.7% | 11.2033 | 12.3826
Si, wt.% | 24.6187 | 0.47999 | 23.6587 | 25.5787 | 23.1787 | 26.0587 | 1.95% | 3.899% | 5.849% | 23.3878 | 25.8497
Sm, ppm | 9.0676 | 0.84593 | 7.3757 | 10.7594 | 6.5298 | 11.6054 | 9.329% | 18.66% | 27.99% | 8.6142 | 9.521
Sr, ppm | 74.9489 | 5.4016 | 64.1457 | 85.7521 | 58.7441 | 91.1537 | 7.207% | 14.41% | 21.62% | 71.2014 | 78.6963
Ta, ppm | 1.9644 | 0.1291 | 1.7062 | 2.2226 | 1.5771 | 2.3517 | 6.572% | 13.14% | 19.72% | 1.8662 | 2.0626
Tb, ppm | 1.1933 | 0.10088 | 0.99153 | 1.3951 | 0.89064 | 1.4959 | 8.454% | 16.91% | 25.36% | 1.1336 | 1.253
Th, ppm | 15.7059 | 0.40445 | 14.897 | 16.5148 | 14.4925 | 16.9192 | 2.575% | 5.15% | 7.725% | 14.9206 | 16.4912
Ti, wt.% | 0.46429 | 0.013518 | 0.43725 | 0.49133 | 0.42374 | 0.50484 | 2.912% | 5.823% | 8.735% | 0.44108 | 0.48751
Tl, ppm | 1.0205 | 0.064571 | 0.89132 | 1.1496 | 0.82675 | 1.2142 | 6.328% | 12.66% | 18.98% | 0.96944 | 1.0715
Tm, ppm | 0.56714 | 0.036203 | 0.49473 | 0.63955 | 0.45853 | 0.67575 | 6.384% | 12.77% | 19.15% | 0.53878 | 0.5955
U, ppm | 9.5113 | 0.47304 | 8.5652 | 10.4573 | 8.0921 | 10.9304 | 4.973% | 9.947% | 14.92% | 9.0357 | 9.9868
V, ppm | 378.7095 | 25.2775 | 328.1546 | 429.2644 | 302.8771 | 454.5419 | 6.675% | 13.35% | 20.02% | 359.774 | 397.645
W, ppm | 2.9367 | 1.0087 | 0.91937 | 4.954 | 0 | 5.9627 | 34.35% | 68.69% | 103% | 2.7899 | 3.0835
Y, ppm | 36.5396 | 2.0317 | 32.4762 | 40.603 | 30.4445 | 42.6347 | 5.56% | 11.12% | 16.68% | 34.7126 | 38.3666
Yb, ppm | 3.4947 | 0.21126 | 3.0722 | 3.9172 | 2.8609 | 4.1285 | 6.045% | 12.09% | 18.14% | 3.32 | 3.6694
Zn, ppm | 31.6644 | 5.084 | 21.4963 | 41.8325 | 16.4123 | 46.9166 | 16.06% | 32.11% | 48.17% | 30.0812 | 33.2476
4-Acid Digestion
Ag, ppm | 0.097963 | 0.014325 | 0.069313 | 0.12661 | 0.054988 | 0.14094 | 14.62% | 29.25% | 43.87% | 0.093065 | 0.10286
Al, wt.% | 6.4043 | 0.14259 | 6.1191 | 6.6894 | 5.9765 | 6.832 | 2.227% | 4.453% | 6.68% | 6.084 | 6.7245
As, ppm | 63.9806 | 3.3712 | 57.2381 | 70.7231 | 53.8669 | 74.0943 | 5.269% | 10.54% | 15.81% | 60.7816 | 67.1796
Ba, ppm | 254.294 | 12.8487 | 228.5966 | 279.9914 | 215.7479 | 292.8401 | 5.053% | 10.11% | 15.16% | 241.5793 | 267.0087
Be, ppm | 4.7558 | 0.25377 | 4.2483 | 5.2633 | 3.9945 | 5.5171 | 5.336% | 10.67% | 16.01% | 4.518 | 4.9936
Bi, ppm | 1.4754 | 0.098348 | 1.2788 | 1.6721 | 1.1804 | 1.7705 | 6.666% | 13.33% | 20% | 1.4017 | 1.5492
Ca, wt.% | 0.14453 | 0.0083828 | 0.12776 | 0.16129 | 0.11938 | 0.16967 | 5.8% | 11.6% | 17.4% | 0.1373 | 0.15175
Cd, ppm | 0.45646 | 0.059499 | 0.33746 | 0.57545 | 0.27796 | 0.63495 | 13.03% | 26.07% | 39.1% | 0.43363 | 0.47928
Ce, ppm | 148.4917 | 6.257 | 135.9777 | 161.0058 | 129.7206 | 167.2628 | 4.214% | 8.427% | 12.64% | 141.0671 | 155.9163
Co, wt.% | 1.102 | 0.036343 | 1.0293 | 1.1747 | 0.99298 | 1.211 | 3.298% | 6.596% | 9.894% | 1.0469 | 1.1571
Cr, ppm | 93.2783 | 8.5732 | 76.1319 | 110.4246 | 67.5587 | 118.9978 | 9.191% | 18.38% | 27.57% | 88.6144 | 97.9422
Cs, ppm | 2.0321 | 0.10367 | 1.8247 | 2.2394 | 1.7211 | 2.3431 | 5.102% | 10.2% | 15.3% | 1.9305 | 2.1337
Cu, wt.% | 4.5807 | 0.13958 | 4.3016 | 4.8599 | 4.162 | 4.9995 | 3.047% | 6.094% | 9.141% | 4.3517 | 4.8098
Dy, ppm | 4.6792 | 0.63754 | 3.4041 | 5.9542 | 2.7666 | 6.5918 | 13.62% | 27.25% | 40.87% | 4.4452 | 4.9131
Er, ppm | 2.7576 | 0.34386 | 2.0698 | 3.4453 | 1.726 | 3.7892 | 12.47% | 24.94% | 37.41% | 2.6197 | 2.8954
Eu, ppm | 0.91493 | 0.044927 | 0.82508 | 1.0048 | 0.78015 | 1.0497 | 4.91% | 9.821% | 14.73% | 0.86919 | 0.96068
Fe, wt.% | 3.8319 | 0.14472 | 3.5425 | 4.1214 | 3.3978 | 4.2661 | 3.777% | 7.553% | 11.33% | 3.6403 | 4.0235
Ga, ppm | 28.3601 | 1.3412 | 25.6778 | 31.0425 | 24.3366 | 32.3837 | 4.729% | 9.458% | 14.19% | 26.9421 | 29.7782
Gd, ppm | 5.1774 | 0.9084 | 3.3606 | 6.9942 | 2.4523 | 7.9026 | 17.55% | 35.09% | 52.64% | 4.9186 | 5.4363
Ge, ppm | 0.19654 | 0.031571 | 0.1334 | 0.25968 | 0.10183 | 0.29125 | 16.06% | 32.13% | 48.19% | 0.18671 | 0.20637
Hf, ppm | 5.3705 | 0.21761 | 4.9352 | 5.8057 | 4.7176 | 6.0233 | 4.052% | 8.104% | 12.16% | 5.1019 | 5.639
Ho, ppm | 0.902 | 0.11027 | 0.68145 | 1.1225 | 0.57117 | 1.2328 | 12.23% | 24.45% | 36.68% | 0.8569 | 0.9471
In, ppm | 2.506 | 0.16504 | 2.1759 | 2.8361 | 2.0109 | 3.0011 | 6.586% | 13.17% | 19.76% | 2.3807 | 2.6313
K, wt.% | 2.5371 | 0.076252 | 2.3846 | 2.6896 | 2.3083 | 2.7658 | 3.006% | 6.011% | 9.017% | 2.4102 | 2.6639
La, ppm | 71.7949 | 5.0579 | 61.6791 | 81.9108 | 56.6211 | 86.9687 | 7.045% | 14.09% | 21.13% | 68.2052 | 75.3847
Li, ppm | 59.3934 | 2.6895 | 54.0144 | 64.7724 | 51.3249 | 67.4618 | 4.528% | 9.057% | 13.58% | 56.4237 | 62.363
Lu, ppm | 0.44151 | 0.026391 | 0.38873 | 0.49429 | 0.36234 | 0.52069 | 5.977% | 11.95% | 17.93% | 0.41944 | 0.46359
Mg, wt.% | 3.0723 | 0.09221 | 2.8879 | 3.2568 | 2.7957 | 3.349 | 3.001% | 6.003% | 9.004% | 2.9187 | 3.226
Mn, wt.% | 0.089023 | 0.0036599 | 0.081704 | 0.096343 | 0.078044 | 0.1 | 4.111% | 8.222% | 12.33% | 0.084572 | 0.093475
Mo, ppm | 17.0696 | 0.9497 | 15.1702 | 18.969 | 14.2205 | 19.9187 | 5.564% | 11.13% | 16.69% | 16.2161 | 17.9231
Na, wt.% | 0.051973 | 0.0067069 | 0.03856 | 0.065387 | 0.031853 | 0.072094 | 12.9% | 25.81% | 38.71% | 0.049375 | 0.054572
Nb, ppm | 17.0885 | 3.8165 | 9.4556 | 24.7215 | 5.6391 | 28.5379 | 22.33% | 44.67% | 67% | 16.2341 | 17.943
Nd, ppm | 64.0192 | 1.9958 | 60.0275 | 68.0109 | 58.0316 | 70.0067 | 3.118% | 6.235% | 9.353% | 60.8182 | 67.2201
Ni, ppm | 220.1353 | 9.0728 | 201.9896 | 238.2809 | 192.9168 | 247.3537 | 4.121% | 8.243% | 12.36% | 209.1285 | 231.142
P, wt.% | 0.034577 | 0.0032014 | 0.028174 | 0.04098 | 0.024973 | 0.044181 | 9.259% | 18.52% | 27.78% | 0.032848 | 0.036306
Pb, ppm | 13.6067 | 1.2522 | 11.1022 | 16.1112 | 9.85 | 17.3634 | 9.203% | 18.41% | 27.61% | 12.9264 | 14.287
Pr, ppm | 16.9339 | 1.3816 | 14.1707 | 19.6971 | 12.789 | 21.0788 | 8.159% | 16.32% | 24.48% | 16.0872 | 17.7806
Rb, ppm | 94.4528 | 4.4199 | 85.613 | 103.2925 | 81.1932 | 107.7124 | 4.679% | 9.359% | 14.04% | 89.7301 | 99.1754
Re, ppm | 0.029874 | 0.0031345 | 0.023605 | 0.036143 | 0.020471 | 0.039278 | 10.49% | 20.98% | 31.48% | 0.02838 | 0.031368
S, wt.% | 4.3937 | 0.16655 | 4.0606 | 4.7268 | 3.8941 | 4.8934 | 3.791% | 7.581% | 11.37% | 4.174 | 4.6134
Sb, ppm | 3.7211 | 0.25177 | 3.2175 | 4.2246 | 2.9658 | 4.4764 | 6.766% | 13.53% | 20.3% | 3.535 | 3.9071
Sc, ppm | 12.5862 | 0.7774 | 11.0314 | 14.141 | 10.254 | 14.9184 | 6.177% | 12.35% | 18.53% | 11.9569 | 13.2155
Se, ppm | 6.9708 | 0.89631 | 5.1782 | 8.7634 | 4.2818 | 9.6597 | 12.86% | 25.72% | 38.57% | 6.6222 | 7.3193
Sm, ppm | 8.8742 | 0.59452 | 7.6852 | 10.0633 | 7.0907 | 10.6578 | 6.699% | 13.4% | 20.1% | 8.4305 | 9.3179
Sn, ppm | 2.9872 | 0.24958 | 2.488 | 3.4863 | 2.2384 | 3.7359 | 8.355% | 16.71% | 25.07% | 2.8378 | 3.1365
Sr, ppm | 69.8607 | 3.6272 | 62.6062 | 77.1152 | 58.9789 | 80.7424 | 5.192% | 10.38% | 15.58% | 66.3677 | 73.3537
Ta, ppm | 1.192 | 0.2354 | 0.72122 | 1.6628 | 0.48582 | 1.8982 | 19.75% | 39.5% | 59.24% | 1.1324 | 1.2516
Tb, ppm | 0.81401 | 0.14789 | 0.51822 | 1.1098 | 0.37033 | 1.2577 | 18.17% | 36.34% | 54.51% | 0.77331 | 0.85471
Te, ppm | 0.059818 | 0.025833 | 0.0081522 | 0.11148 | 0 | 0.13732 | 43.19% | 86.37% | 129.6% | 0.056827 | 0.062809
Th, ppm | 15.977 | 0.95834 | 14.0603 | 17.8937 | 13.102 | 18.852 | 5.998% | 12% | 17.99% | 15.1781 | 16.7758
Ti, wt.% | 0.27375 | 0.080132 | 0.11349 | 0.43402 | 0.033354 | 0.51415 | 29.27% | 58.54% | 87.82% | 0.26006 | 0.28744
Tl, ppm | 0.9867 | 0.068305 | 0.85009 | 1.1233 | 0.78179 | 1.1916 | 6.923% | 13.85% | 20.77% | 0.93736 | 1.036
Tm, ppm | 0.41895 | 0.059903 | 0.29915 | 0.53876 | 0.23924 | 0.59866 | 14.3% | 28.6% | 42.89% | 0.39801 | 0.4399
U, ppm | 9.2368 | 0.71151 | 7.8138 | 10.6598 | 7.1023 | 11.3713 | 7.703% | 15.41% | 23.11% | 8.775 | 9.6986
V, ppm | 356.6079 | 15.7495 | 325.109 | 388.1069 | 309.3595 | 403.8564 | 4.416% | 8.833% | 13.25% | 338.7775 | 374.4383
W, ppm | 2.7793 | 0.25419 | 2.2709 | 3.2877 | 2.0168 | 3.5419 | 9.146% | 18.29% | 27.44% | 2.6404 | 2.9183
Y, ppm | 24.4386 | 1.0002 | 22.4382 | 26.439 | 21.438 | 27.4392 | 4.093% | 8.186% | 12.28% | 23.2167 | 25.6605
Yb, ppm | 2.8839 | 0.31429 | 2.2553 | 3.5125 | 1.941 | 3.8268 | 10.9% | 21.8% | 32.69% | 2.7397 | 3.0281
Zn, ppm | 29.0115 | 2.0849 | 24.8417 | 33.1812 | 22.7568 | 35.2661 | 7.186% | 14.37% | 21.56% | 27.5609 | 30.462
Zr, ppm | 187.9751 | 6.3881 | 175.1989 | 200.7513 | 168.8108 | 207.1394 | 3.398% | 6.797% | 10.2% | 178.5763 | 197.3739
Aqua Regia Digestion
Ag, ppm | 0.075622 | 0.012766 | 0.050089 | 0.10115 | 0.037323 | 0.11392 | 16.88% | 33.76% | 50.65% | 0.071841 | 0.079403
Al, wt.% | 0.9585 | 0.18674 | 0.58503 | 1.332 | 0.39829 | 1.5187 | 19.48% | 38.96% | 58.45% | 0.91058 | 1.0064
As, ppm | 61.6015 | 3.2077 | 55.1862 | 68.0168 | 51.9785 | 71.2245 | 5.207% | 10.41% | 15.62% | 58.5214 | 64.6816
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 56.6609 | 8.5721 | 39.5166 | 73.8052 | 30.9445 | 82.3773 | 15.13% | 30.26% | 45.39% | 53.8279 | 59.494
Be, ppm | 2.2126 | 0.27077 | 1.6711 | 2.7542 | 1.4003 | 3.0249 | 12.24% | 24.48% | 36.71% | 2.102 | 2.3232
Bi, ppm | 1.3874 | 0.077887 | 1.2316 | 1.5431 | 1.1537 | 1.621 | 5.614% | 11.23% | 16.84% | 1.318 | 1.4567
Ca, wt.% | 0.13664 | 0.0072653 | 0.12211 | 0.15117 | 0.11484 | 0.15844 | 5.317% | 10.63% | 15.95% | 0.12981 | 0.14347
Cd, ppm | 0.46427 | 0.025114 | 0.41405 | 0.5145 | 0.38893 | 0.53962 | 5.409% | 10.82% | 16.23% | 0.44106 | 0.48749
Ce, ppm | 43.6838 | 7.9714 | 27.741 | 59.6266 | 19.7696 | 67.5981 | 18.25% | 36.5% | 54.74% | 41.4996 | 45.868
Co, wt.% | 1.0756 | 0.038571 | 0.99849 | 1.1528 | 0.95991 | 1.1913 | 3.586% | 7.172% | 10.76% | 1.0218 | 1.1294
Cr, ppm | 35.2455 | 3.7553 | 27.735 | 42.756 | 23.9797 | 46.5113 | 10.65% | 21.31% | 31.96% | 33.4832 | 37.0078
Cs, ppm | 0.32386 | 0.10366 | 0.11654 | 0.53118 | 0.012883 | 0.63483 | 32.01% | 64.01% | 96.02% | 0.30767 | 0.34005
Cu, wt.% | 4.5716 | 0.10799 | 4.3556 | 4.7875 | 4.2476 | 4.8955 | 2.362% | 4.724% | 7.087% | 4.343 | 4.8001
Dy, ppm | 0.99577 | 0.10301 | 0.78975 | 1.2018 | 0.68674 | 1.3048 | 10.34% | 20.69% | 31.03% | 0.94598 | 1.0456
Er, ppm | 0.4069 | 0.081089 | 0.24472 | 0.56907 | 0.16363 | 0.65016 | 19.93% | 39.86% | 59.79% | 0.38655 | 0.42724
Eu, ppm | 0.38641 | 0.034922 | 0.31657 | 0.45626 | 0.28165 | 0.49118 | 9.037% | 18.07% | 27.11% | 0.36709 | 0.40573
Fe, wt.% | 3.6635 | 0.17838 | 3.3067 | 4.0202 | 3.1283 | 4.1986 | 4.869% | 9.738% | 14.61% | 3.4803 | 3.8466
Ga, ppm | 8.2218 | 1.4904 | 5.2409 | 11.2027 | 3.7505 | 12.6931 | 18.13% | 36.26% | 54.38% | 7.8107 | 8.6329
Gd, ppm | 2.025 | 0.13973 | 1.7456 | 2.3045 | 1.6058 | 2.4442 | 6.9% | 13.8% | 20.7% | 1.9238 | 2.1263
Ge, ppm | 0.095476 | 0.027455 | 0.040566 | 0.15039 | 0.013111 | 0.17784 | 28.76% | 57.51% | 86.27% | 0.090702 | 0.10025
Hf, ppm | 0.33039 | 0.046972 | 0.23644 | 0.42433 | 0.18947 | 0.4713 | 14.22% | 28.43% | 42.65% | 0.31387 | 0.3469
Hg, ppm | 0.060264 | 0.011797 | 0.036669 | 0.083858 | 0.024872 | 0.095655 | 19.58% | 39.15% | 58.73% | 0.057251 | 0.063277
Ho, ppm | 0.16772 | 0.02755 | 0.11262 | 0.22281 | 0.085066 | 0.25036 | 16.43% | 32.85% | 49.28% | 0.15933 | 0.1761
In, ppm | 1.827 | 0.10922 | 1.6085 | 2.0454 | 1.4993 | 2.1546 | 5.978% | 11.96% | 17.93% | 1.7356 | 1.9183
K, wt.% | 0.19566 | 0.029826 | 0.13601 | 0.25531 | 0.10619 | 0.28514 | 15.24% | 30.49% | 45.73% | 0.18588 | 0.20545
La, ppm | 20.3761 | 3.1929 | 13.9903 | 26.762 | 10.7974 | 29.9549 | 15.67% | 31.34% | 47.01% | 19.3573 | 21.395
Li, ppm | 35.3459 | 7.3347 | 20.6764 | 50.0153 | 13.3417 | 57.35 | 20.75% | 41.5% | 62.25% | 33.5786 | 37.1132
Mg, wt.% | 2.4173 | 0.23861 | 1.94 | 2.8945 | 1.7014 | 3.1331 | 9.871% | 19.74% | 29.61% | 2.2964 | 2.5381
Mn, wt.% | 0.088654 | 0.0040845 | 0.080485 | 0.096823 | 0.076401 | 0.10091 | 4.607% | 9.214% | 13.82% | 0.084221 | 0.093087
Mo, ppm | 16.2231 | 0.90964 | 14.4038 | 18.0423 | 13.4941 | 18.952 | 5.607% | 11.21% | 16.82% | 15.4119 | 17.0342
Na, wt.% | 0.010713 | 0.0022649 | 0.0061836 | 0.015243 | 0.0039187 | 0.017508 | 21.14% | 42.28% | 63.42% | 0.010178 | 0.011249
Nb, ppm | 0.093611 | 0.024008 | 0.045595 | 0.14163 | 0.021587 | 0.16564 | 25.65% | 51.29% | 76.94% | 0.088931 | 0.098292
Nd, ppm | 22.9844 | 3.0653 | 16.8539 | 29.115 | 13.7886 | 32.1802 | 13.34% | 26.67% | 40.01% | 21.8352 | 24.1336
Ni, ppm | 213.8519 | 12.5678 | 188.7164 | 238.9874 | 176.1486 | 251.5551 | 5.877% | 11.75% | 17.63% | 203.1593 | 224.5444
P, wt.% | 0.020017 | 0.0011353 | 0.017746 | 0.022288 | 0.016611 | 0.023423 | 5.672% | 11.34% | 17.02% | 0.019016 | 0.021018
Pb, ppm | 10.9141 | 1.0793 | 8.7555 | 13.0726 | 7.6762 | 14.1519 | 9.889% | 19.78% | 29.67% | 10.3684 | 11.4598
Pr, ppm | 5.4327 | 1.4331 | 2.5666 | 8.2988 | 1.1335 | 9.7318 | 26.38% | 52.76% | 79.14% | 5.161 | 5.7043
Rb, ppm | 6.4953 | 1.1161 | 4.2631 | 8.7275 | 3.147 | 9.8436 | 17.18% | 34.37% | 51.55% | 6.1706 | 6.8201
Re, ppm | 0.026561 | 0.0027215 | 0.021118 | 0.032004 | 0.018396 | 0.034726 | 10.25% | 20.49% | 30.74% | 0.025233 | 0.027889
S, wt.% | 4.3112 | 0.1697 | 3.9718 | 4.6507 | 3.8021 | 4.8204 | 3.936% | 7.873% | 11.81% | 4.0957 | 4.5268
Sb, ppm | 2.1662 | 0.40319 | 1.3598 | 2.9726 | 0.95663 | 3.3758 | 18.61% | 37.23% | 55.84% | 2.0579 | 2.2745
Sc, ppm | 3.2952 | 0.34606 | 2.6031 | 3.9873 | 2.257 | 4.3334 | 10.5% | 21% | 31.51% | 3.1304 | 3.46
Se, ppm | 6.6992 | 0.52834 | 5.6425 | 7.7558 | 5.1141 | 8.2842 | 7.887% | 15.77% | 23.66% | 6.3642 | 7.0341
Sm, ppm | 3.5079 | 0.47323 | 2.5614 | 4.4543 | 2.0882 | 4.9276 | 13.49% | 26.98% | 40.47% | 3.3325 | 3.6833
Sn, ppm | 1.5356 | 0.15227 | 1.2311 | 1.8402 | 1.0788 | 1.9924 | 9.916% | 19.83% | 29.75% | 1.4588 | 1.6124
Sr, ppm | 15.1067 | 2.6228 | 9.861 | 20.3523 | 7.2382 | 22.9751 | 17.36% | 34.72% | 52.09% | 14.3513 | 15.862
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.20531 | 0.03494 | 0.13543 | 0.27519 | 0.10049 | 0.31013 | 17.02% | 34.04% | 51.05% | 0.19504 | 0.21558
Te, ppm | 0.053516 | 0.012948 | 0.02762 | 0.079411 | 0.014672 | 0.092359 | 24.19% | 48.39% | 72.58% | 0.05084 | 0.056191
Th, ppm | 6.3228 | 0.9101 | 4.5026 | 8.143 | 3.5925 | 9.0531 | 14.39% | 28.79% | 43.18% | 6.0066 | 6.6389
U, ppm | 3.491 | 0.27995 | 2.9311 | 4.0509 | 2.6511 | 4.3308 | 8.019% | 16.04% | 24.06% | 3.3164 | 3.6655
V, ppm | 62.0724 | 8.7136 | 44.6452 | 79.4996 | 35.9316 | 88.2132 | 14.04% | 28.08% | 42.11% | 58.9688 | 65.176
W, ppm | 1.0849 | 0.14548 | 0.79391 | 1.3758 | 0.64843 | 1.5213 | 13.41% | 26.82% | 40.23% | 1.0306 | 1.1391
Y, ppm | 3.9366 | 0.48682 | 2.9629 | 4.9102 | 2.4761 | 5.397 | 12.37% | 24.73% | 37.1% | 3.7397 | 4.1334
Yb, ppm | 0.31535 | 0.07399 | 0.16737 | 0.46333 | 0.093377 | 0.53732 | 23.46% | 46.93% | 70.39% | 0.29958 | 0.33112
Zn, ppm | 26.6028 | 1.2871 | 24.0287 | 29.1769 | 22.7416 | 30.464 | 4.838% | 9.676% | 14.51% | 25.2727 | 27.933
Zr, ppm | 10.1145 | 1.3274 | 7.4597 | 12.7693 | 6.1323 | 14.0967 | 13.12% | 26.25% | 39.37% | 9.6088 | 10.6202
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 555b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AL-5S*AAS | 5% sulphuric acid laech with atomic absorption spectroscopy
AL-5S*OES | 5% sulphuric acid leach with inductively coupled plasma optical emission spectroscopy
AL-5S*OES/AAS | 5% sulphuric acid leach with AAS or ICP-OES finish
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PPP*XRF | pressed powder pellet with X-ray fluorescence spectroscopy
Text Values:
NA | Unable to report due to QC failure (Lab 19)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 555b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Labwest Minerals Analysis, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Fusion XRF
| Analytical results for Ag in OREAS 555b (Indicative Value 0.035 ppm) | [21 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 3
3
| 1 | 1 | < 100 | 0.022 | [18 empty] | 1
| 1 | 2 | < 100 | 0.044 | [18 empty] | 1
| 1 | 3 | < 100 | 0.02 | [18 empty] | 16
| 1 | 4 | < 100 | 0.046 | [18 empty] | 0.0345
| 1 | 5 | < 100 | 0.028 | [18 empty] | 7
| 1 | 6 | < 100 | 0.047
| Mean |  |  | 0.0345
| Median |  |  | 0.036
| Std Dev. |  |  | 0.01255
| Rel.Std.Dev. |  |  | 36.38%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Al2O3 in OREAS 555b (Certified Value 12.39 wt.%) | [21 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 13.68 | 12.53 | 12.49 | 12.56 | 12.17 | 12.22 | 12.5 | 12.4 | 12.5 | 12.59 | 12.3521 | 10.8 | 12.19 | 13.31 | 12.5 | 12.024 | 11.19 | 12.63 | 12.35 |  | 1
| 1 | 2 | 13.82 | 12.41 | 12.44 | 12.57 | 12.05 | 12.21 | 12.55 | 12.4 | 12.7 | 12.56 | 12.2861 | 10.7 | 12.3 | 13.3 | 12.5 | 12.1109 | 11.8 | 12.5 | 12.29 |  | 29
| 1 | 3 | 13.76 | 12.38 | 12.49 | 12.52 | 12.09 | 12.22 | 12.55 | 12.4 | 12.8 | 12.67 | 12.3693 | 10.8 | 12.26 | 13 | 12.5 | 12.0483 | 11.36 | 12.48 | 12.32 |  | 16
| 1 | 4 | 13.74 | 12.48 | 12.41 | 12.49 | 12.11 | 12.21 | 12.35 | 12.4 | 12.75 | 12.54 | 12.3875 | 10.8 | 12.33 | 13.28 | 12.5 | 12.0695 | 12.63 | 12.49 | 12.34 |  | 12.3891
| 1 | 5 | 13.79 | 12.46 | 12.46 | 12.53 | 12 | 12.21 | 12.55 | 12.4 | 12.5 | 12.67 | 12.4038 | 10.8 | 12.33 | 13.13 | 12.5 | 12.0846 | 13.38 | 12.53 | 12.32 |  | 7
| 1 | 6 | 13.67 | 12.42 | 12.46 | 12.41 | 12.16 | 12.29 | 12.45 | 12.4 | 12.65 | 12.54 | 12.2811 | 10.8 | 12.3 | 12.99 | 12.4 | 12.0645 | 12.5 | 12.54 | 12.34
| Mean |  | 13.7433 | 12.4467 | 12.4583 | 12.5133 | 12.0967 | 12.2267 | 12.4917 | 12.4 | 12.65 | 12.595 | 12.3466 | 10.7833 | 12.285 | 13.1683 | 12.4833 | 12.067 | 12.1433 | 12.5283 | 12.3267
| Median |  | 13.75 | 12.44 | 12.46 | 12.525 | 12.1 | 12.215 | 12.525 | 12.4 | 12.675 | 12.575 | 12.3607 | 10.8 | 12.3 | 13.205 | 12.5 | 12.067 | 12.15 | 12.515 | 12.33
| Std Dev. |  | 0.059554 | 0.054283 | 0.030605 | 0.058195 | 0.065013 | 0.031411 | 0.080104 | 0 | 0.12649 | 0.06091 | 0.051844 | 0.040825 | 0.053198 | 0.14932 | 0.040825 | 0.029823 | 0.84068 | 0.054924 | 0.021602
| Rel.Std.Dev. |  | 0.4333% | 0.4361% | 0.2457% | 0.4651% | 0.5374% | 0.2569% | 0.6413% | 0% | 0.9999% | 0.4836% | 0.4199% | 0.3786% | 0.433% | 1.134% | 0.327% | 0.2471% | 6.923% | 0.4384% | 0.1752%
| PDM3 |  | 10.93% | 0.4648% | 0.559% | 1.003% | -2.36% | -1.311% | 0.828% | 0.08814% | 2.106% | 1.662% | -0.3425% | -12.96% | -0.8401% | 6.29% | 0.7608% | -2.6% | -1.984% | 1.124% | -0.5038%
| Z-Score (Absolute) |  | 5.89 | 0 | 0.05 | 0.3 | 1.59 | 1 | 0.2 | 0.21 | 0.92 | 0.67 | 0.45 | 7.56 | 0.73 | 3.28 | 0.17 | 1.73 | 1.38 | 0.37 | 0.55
| Analytical results for As in OREAS 555b (Indicative Value 60 ppm) | [21 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 08 | 12 | 13 | 17 | 24 | 25 | 28 | [12 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [12 empty] | 0
0
| 1 | 1 | 100 | 100 | < 100 | < 100 | 70 | 45.4442 | 60 | < 100 | [12 empty] | 1
| 1 | 2 | 100 | < 100 | < 100 | < 100 | 70 | 53.0183 | 60 | < 100 | [12 empty] | 2
| 1 | 3 | 100 | < 100 | 100 | < 100 | 70 | 45.4442 | 60 | < 100 | [12 empty] | 16
| 1 | 4 | 100 | < 100 | 100 | < 100 | 70 | 45.4442 | 61 | < 100 | [12 empty] | 60.2222
| 1 | 5 | 100 | < 100 | < 100 | < 100 | 70 | 45.4442 | 59 | < 100 | [12 empty] | 8
| 1 | 6 | 100 | < 100 | < 100 | < 100 | 70 | 45.4442 | 64 | < 100
| Mean |  | 100 | 100 | 100 |  | 70 | 46.7066 | 60.6667
| Median |  | 100 | 100 | 100 |  | 70 | 45.4442 | 60
| Std Dev. |  | 0 |  | 0 |  | 0 | 3.0921 | 1.7512
| Rel.Std.Dev. |  | 0% |  | 0% |  | 0% | 6.62% | 2.887%
| PDM3 |  | 66.05% | 66.05% | 66.05% |  | 16.24% | -22.44% | 0.7381%
| Z-Score (Absolute) |  | 2.87 | 0.08 | 0.51 | 0.67 | 0.75 | 0.91 | 0.08 | 0.67
| Analytical results for BaO in OREAS 555b (Certified Value 316 ppm) | [21 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 28 |  |  |  |  | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | 500 | 300 | 310 | 300 | 200 | 300 | 440 | 378.88 | < 1000 | 292.5 | 400 | 300 | 283.05 | 260.1 | 110 |  |  |  |  | 1
| 1 | 2 | < 100 | 300 | 300 | 300 | 300 | 200 | 300 | 420 | 337.92 | < 1000 | 299.2 | 400 | 300 | 306 | 253.4 | 110 |  |  |  |  | #N/A
| 1 | 3 | < 100 | 400 | 400 | 340 | 300 | 200 | 300 | 460 | 409.6 | < 1000 | 297 | 400 | 300 | 336.6 | 263.5 | 110 |  |  |  |  | 16
| 1 | 4 | < 100 | 500 | 400 | 330 | 300 | 200 | 300 | 480 | 337.92 | < 1000 | 292.5 | 400 | 300 | 285.6 | 260.1 | 110 |  |  |  |  | 316.1379
| 1 | 5 | < 100 | 500 | 300 | 320 | 300 | 300 | 300 | 470 | 348.16 | < 1000 | 289.2 | 400 | 300 | 275.4 | 260.1 | < 110 |  |  |  |  | 8
| 1 | 6 | < 100 | 200 | 300 | 320 | 300 | 300 | 300 | 450 | 378.88 | < 1000 | 290.3 | 300 | 300 | 346.8 | 259 | < 110
| Mean |  |  | 400 | 333.3333 | 320 | 300 | 233.3333 | 300 | 453.3333 | 365.2267 |  | 293.45 | 383.3333 | 300 | 305.575 | 259.3667 | 110
| Median |  |  | 450 | 300 | 320 | 300 | 200 | 300 | 455 | 363.52 |  | 292.5 | 400 | 300 | 295.8 | 260.1 | 110
| Std Dev. |  |  | 126.4911 | 51.6398 | 14.1421 | 0 | 51.6398 | 0 | 21.6025 | 28.7207 |  | 3.8847 | 40.8248 | 0 | 29.9304 | 3.2995 | 0
| Rel.Std.Dev. |  |  | 31.62% | 15.49% | 4.419% | 0% | 22.13% | 0% | 4.765% | 7.864% |  | 1.324% | 10.65% | 0% | 9.795% | 1.272% | 0%
| PDM3 |  |  | 26.53% | 5.439% | 1.222% | -5.105% | -26.19% | -5.105% | 43.4% | 15.53% |  | -7.177% | 21.26% | -5.105% | -3.341% | -17.96% | -65.21%
| Z-Score (Absolute) |  | 3.22 | 1.24 | 0.39 | 0.22 | 0.04 | 0.89 | 0.04 | 1.92 | 0.8 | 2.51 | 0.12 | 1.03 | 0.04 | 0.04 | 0.55 | 2.67
| Analytical results for Be in OREAS 555b (Indicative Value 3 ppm) | [21 empty] | Indicative
Be | Round | Replicate | Lab | [19 empty] | 1
| No | No | 25 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | 3 | [19 empty] | 1
| 1 | 2 | 3 | [19 empty] | 3
| 1 | 3 | 3 | [19 empty] | 16
| 1 | 4 | 3 | [19 empty] | 3
| 1 | 5 | 3 | [19 empty] | 9
| 1 | 6 | 3
| Mean |  | 3
| Median |  | 3
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 555b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 08 | 12 | 13 | 25 | [15 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | 0.9 | [15 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | 0.8 | [15 empty] | 4
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | 0.9 | [15 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | 1.1 | [15 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | 1.1 | [15 empty] | 10
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | 1.4
| Mean | [5 empty] | 1.0333
| Median | [5 empty] | 1
| Std Dev. | [5 empty] | 0.21602
| Rel.Std.Dev. | [5 empty] | 20.91%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 555b (Certified Value 0.201 wt.%) | [21 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.24 | 0.15 | 0.2 | 0.2 | 0.19 | 0.2 | 0.21 | 0.19 | 0.2 | 0.14 | 0.19306 | 0.1 | 0.17 | 0.21 | 0.21 | 0.18988 | 0.23 | 0.2 | 0.2 |  | 1
| 1 | 2 | 0.23 | 0.15 | 0.2 | 0.2 | 0.19 | 0.203 | 0.21 | 0.2 | 0.21 | 0.14 | 0.19457 | 0.2 | 0.18 | 0.21 | 0.21 | 0.19897 | 0.15 | 0.2 | 0.2 |  | 33
| 1 | 3 | 0.23 | 0.15 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21 | 0.2 | 0.21 | 0.15 | 0.19959 | 0.2 | 0.18 | 0.2 | 0.2 | 0.19897 | 0.15 | 0.2 | 0.2 |  | 16
| 1 | 4 | 0.23 | 0.15 | 0.2 | 0.21 | 0.2 | 0.197 | 0.2 | 0.2 | 0.21 | 0.14 | 0.19657 | 0.2 | 0.18 | 0.22 | 0.21 | 0.1919 | 0.16 | 0.2 | 0.2 |  | 0.20149
| 1 | 5 | 0.24 | 0.15 | 0.2 | 0.21 | 0.2 | 0.202 | 0.21 | 0.2 | 0.2 | 0.14 | 0.19657 | 0.2 | 0.18 | 0.22 | 0.21 | 0.19089 | 0.18 | 0.2 | 0.2 |  | 9
| 1 | 6 | 0.25 | 0.14 | 0.2 | 0.2 | 0.2 | 0.195 | 0.21 | 0.2 | 0.2 | 0.14 | 0.19657 | 0.2 | 0.17 | 0.21 | 0.21 | 0.1919 | 0.18 | 0.2 | 0.19
| Mean |  | 0.23667 | 0.14833 | 0.2 | 0.20333 | 0.19667 | 0.1995 | 0.20833 | 0.19833 | 0.205 | 0.14167 | 0.19616 | 0.18333 | 0.17667 | 0.21167 | 0.20833 | 0.19375 | 0.175 | 0.2 | 0.19833
| Median |  | 0.235 | 0.15 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21 | 0.2 | 0.205 | 0.14 | 0.19657 | 0.2 | 0.18 | 0.21 | 0.21 | 0.1919 | 0.17 | 0.2 | 0.2
| Std Dev. |  | 0.008165 | 0.0040825 | 3.0405e-17 | 0.005164 | 0.005164 | 0.0030166 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0040825 | 0.002209 | 0.040825 | 0.005164 | 0.0075277 | 0.0040825 | 0.0041109 | 0.030166 | 3.0405e-17 | 0.0040825
| Rel.Std.Dev. |  | 3.45% | 2.752% | 1.52e-14% | 2.54% | 2.626% | 1.512% | 1.96% | 2.058% | 2.672% | 2.882% | 1.126% | 22.27% | 2.923% | 3.556% | 1.96% | 2.122% | 17.24% | 1.52e-14% | 2.058%
| PDM3 |  | 17.46% | -26.38% | -0.7409% | 0.9134% | -2.395% | -0.989% | 3.395% | -1.568% | 1.741% | -29.69% | -2.649% | -9.012% | -12.32% | 5.049% | 3.395% | -3.842% | -13.15% | -0.7409% | -1.568%
| Z-Score (Absolute) |  | 4.53 | 6.06 | 0.13 | 0.53 | 0.27 | 0.07 | 1.13 | 0.07 | 0.73 | 6.86 | 0.33 | NA | 2.67 | 1.53 | 1.13 | 0.62 | 2.87 | 0.13 | 0.07
Results from laboratory 19 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 555b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 13 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | < 100 | [18 empty] | 5
| 1 | 3 | < 100 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | < 100 | [18 empty] | 11
| 1 | 6 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 555b (Indicative Value 130 ppm) | [21 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 12 | 13 | 25 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 80 | 80 | 149.014 | 153 | [16 empty] | 1
| 1 | 2 | 80 | 160 | 148.535 | 158 | [16 empty] | 6
| 1 | 3 | 80 | 160 | 148.687 | 157 | [16 empty] | 16
| 1 | 4 | 80 | 160 | 148.734 | 153 | [16 empty] | 129.9572
| 1 | 5 | 80 | 160 | 148.634 | 149 | [16 empty] | 12
| 1 | 6 | 80 | 80 | 148.862 | 154
| Mean |  | 80 | 133.3333 | 148.7443 | 154
| Median |  | 80 | 160 | 148.7105 | 153.5
| Std Dev. |  | 0 | 41.3118 | 0.17085 | 3.2249
| Rel.Std.Dev. |  | 0% | 30.98% | 0.1149% | 2.094%
| PDM3 |  | -38.44% | 2.598% | 14.46% | 18.5%
| Z-Score (Absolute) |  | 4.48 | 0.48 | 0.48 | 0.87
| Analytical results for Cl in OREAS 555b (Indicative Value 224 ppm) | [21 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 17 | 24 | 28 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [17 empty] | 0
0
| 1 | 1 | 210 | 320 | 100 | [17 empty] | 1
| 1 | 2 | 210 | 351.1111 | 100 | [17 empty] | 7
| 1 | 3 | 210 | 375 | 100 | [17 empty] | 16
| 1 | 4 | 220 | 370 | 100 | [17 empty] | 223.8765
| 1 | 5 | 210 | 351.6667 | 100 | [17 empty] | 13
| 1 | 6 | 240 | 390 | 100
| Mean |  | 216.6667 | 359.6296 | 100
| Median |  | 210 | 360.8333 | 100
| Std Dev. |  | 12.1106 | 24.3829 | 0
| Rel.Std.Dev. |  | 5.59% | 6.78% | 0%
| PDM3 |  | -3.22% | 60.64% | -55.33%
| Z-Score (Absolute) |  | 0 | 0.83 | 0.67
| Analytical results for Co in OREAS 555b (Certified Value 1.08 wt.%) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 17 | 19 | 22 | 23 | 24 | 25 | 27 | 28 |  |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 1.09 | 1.13 | 1.106 | 1.09 | 1.0787 | 1.11 | 1.08 | 1.09 | 1.07 | 1.1839 | 1.198 | 1 | 1.144 | 0.896 | 1.0511 | 1.06 | 1.094 | 1.18 |  |  | 1
| 1 | 2 | 1.09 | 1.14 | 1.09 | 1.08 | 1.0781 | 1.07 | 1.12 | 1.09 | 1.08 | 1.1627 | 1.192 | 1.01 | 1.1574 | 0.888 | 1.088 | 1.07 | 1.073 | 1.17 |  |  | 36
| 1 | 3 | 1.09 | 1.13 | 1.109 | 1.09 | 1.0839 | 1.08 | 1.08 | 1.06 | 1.08 | 1.1813 | 1.196 | 1.02 | 1.1548 | 0.896 | 1.0547 | 1.1 | 1.095 | 1.19 |  |  | 16
| 1 | 4 | 1.09 | 1.13 | 1.112 | 1.09 | 1.1125 | 1.06 | 1.08 | 1.09 | 1.08 | 1.1803 | 1.199 | 1.01 | 1.1684 | 0.888 | 1.0563 | 1.04 | 1.073 | 1.18 |  |  | 1.0848
| 1 | 5 | 1.09 | 1.14 | 1.107 | 1.09 | 1.0959 | 1.08 | 1.07 | 1.09 | 1.06 | 1.1814 | 1.208 | 1.01 | 1.1537 | 0.881 | 1.0282 | 1.04 | 1.081 | 1.19 |  |  | 10
| 1 | 6 | 1.09 | 1.14 | 1.11 | 1.08 | 1.1027 | 1.06 | 1.08 | 1.08 | 1.08 | 1.1815 | 1.204 | 1 | 1.1523 | 0.896 | 1.0732 | 1.01 | 1.082 | 1.18
| Mean |  | 1.09 | 1.135 | 1.1057 | 1.0867 | 1.092 | 1.0767 | 1.085 | 1.0833 | 1.075 | 1.1785 | 1.1995 | 1.0083 | 1.1551 | 0.89083 | 1.0586 | 1.0533 | 1.083 | 1.1817
| Median |  | 1.09 | 1.135 | 1.108 | 1.09 | 1.0899 | 1.075 | 1.08 | 1.09 | 1.08 | 1.1814 | 1.1985 | 1.01 | 1.1542 | 0.892 | 1.0555 | 1.05 | 1.0815 | 1.18
| Std Dev. |  | 0 | 0.0054772 | 0.0079666 | 0.005164 | 0.01404 | 0.018619 | 0.017607 | 0.012111 | 0.0083666 | 0.0078395 | 0.0057184 | 0.0075277 | 0.0079408 | 0.0062102 | 0.0204 | 0.030768 | 0.0096954 | 0.0075277
| Rel.Std.Dev. |  | 0% | 0.4826% | 0.7205% | 0.4752% | 1.286% | 1.729% | 1.623% | 1.118% | 0.7783% | 0.6652% | 0.4767% | 0.7466% | 0.6875% | 0.6971% | 1.927% | 2.921% | 0.8952% | 0.637%
| PDM3 |  | 0.4764% | 4.625% | 1.921% | 0.1691% | 0.6577% | -0.7527% | 0.0155% | -0.1381% | -0.9063% | 8.636% | 10.57% | -7.052% | 6.477% | -17.88% | -2.42% | -2.904% | -0.1689% | 8.926%
| Z-Score (Absolute) |  | 0.12 | 1.41 | 0.57 | 0.02 | 0.18 | 0.26 | 0.02 | 0.07 | 0.31 | 2.65 | 3.26 | 2.22 | 1.98 | 5.58 | 0.78 | 0.93 | 0.08 | 2.74
| Analytical results for Cr2O3 in OREAS 555b (Indicative Value 159 ppm) | [21 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 25 | 27 | 28 |  |  |  | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | 100 | 300 | 200 | 170 | 150 | 100 | 150 | 130 | 210 | < 1000 | 200 | 200 | 500 | 141 | 160 | 200 |  |  |  | 1
| 1 | 2 | < 100 | 90000 | 200 | 200 | 140 | 150 | 100 | 150 | < 50 | 220 | < 1000 | 200 | 200 | 600 | 132 | 160 | 200 |  |  |  | 8
| 1 | 3 | < 100 | 100 | 200 | 100 | 160 | 150 | 100 | 290 | < 50 | 220 | < 1000 | 200 | 200 | 600 | 157 | 150 | 100 |  |  |  | 16
| 1 | 4 | < 100 | 100 | 200 | 200 | 170 | 150 | 100 | 150 | 50 | 200 | < 1000 | 200 | 200 | 500 | 144 | 150 | 100 |  |  |  | 159.3327
| 1 | 5 | < 100 | 100 | 200 | 200 | 150 | 150 | 100 | 150 | < 50 | 240 | < 1000 | 200 | 200 | 500 | 155 | 180 | 100 |  |  |  | 14
| 1 | 6 | < 100 | 100 | 400 | 200 | 160 | 150 | 100 | 150 | < 50 | 230 | < 1000 | 200 | 200 | 500 | 175 | 150 | 100
| Mean |  |  | 15083.3333 | 250 | 183.3333 | 158.3333 | 150 | 100 | 173.3333 | 90 | 220 |  | 200 | 200 | 533.3333 | 150.6667 | 158.3333 | 133.3333
| Median |  |  | 100 | 200 | 200 | 160 | 150 | 100 | 150 | 90 | 220 |  | 200 | 200 | 500 | 149.5 | 155 | 100
| Std Dev. |  |  | 36701.5213 | 83.666 | 40.8248 | 11.6905 | 0 | 0 | 57.1548 | 56.5685 | 14.1421 |  | 0 | 0 | 51.6398 | 15.082 | 11.6905 | 51.6398
| Rel.Std.Dev. |  |  | 243.3% | 33.47% | 22.27% | 7.383% | 0% | 0% | 32.97% | 62.85% | 6.428% |  | 0% | 0% | 9.682% | 10.01% | 7.383% | 38.73%
| PDM3 |  |  | 9367% | 56.9% | 15.06% | -0.6272% | -5.857% | -37.24% | 8.787% | -43.51% | 38.08% |  | 25.52% | 25.52% | 234.7% | -5.439% | -0.6272% | -16.32%
| Z-Score (Absolute) |  | 2.19 | 270.5 | 1.44 | 0.23 | 0.22 | 0.44 | 1.28 | 0 | 2.25 | 0.9 | 5.98 | 0.53 | 0.53 | 6.58 | 0.36 | 0.22 | 0.67
| Analytical results for Cs in OREAS 555b (Indicative Value 2.22 ppm) | [21 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | < 100 | 2.2 | [18 empty] | 1
| 1 | 2 | < 100 | 2.2 | [18 empty] | 9
| 1 | 3 | < 100 | 2.3 | [18 empty] | 16
| 1 | 4 | < 100 | 2.2 | [18 empty] | 2.2167
| 1 | 5 | < 100 | 2.2 | [18 empty] | 15
| 1 | 6 | < 100 | 2.2
| Mean |  |  | 2.2167
| Median |  |  | 2.2
| Std Dev. |  |  | 0.040825
| Rel.Std.Dev. |  |  | 1.842%
| PDM3 |  |  | -1.443e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 555b (Certified Value 4.57 wt.%) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 17 | 19 | 22 | 23 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 4.48 | 4.77 | 4.671 | 4.5854 | 4.37 | 4.7005 | 4.66 | 4.59 | 4.55 | 4.52 | 4.3229 | > 5 | 4.234 | 4.673 | 4.38 | 4.3827 | 4.89 | 4.585 | 4.84 |  | 1
| 1 | 2 | 4.47 | 4.75 | 4.672 | 4.5295 | 4.362 | 4.6397 | 4.66 | 4.55 | 4.6 | 4.65 | 4.3562 | > 5 | 4.245 | 4.785 | 4.37 | 4.5075 | 4.7 | 4.517 | 4.77 |  | 39
| 1 | 3 | 4.48 | 4.8 | 4.67 | 4.5774 | 4.314 | 4.676 | 4.65 | 4.56 | 4.56 | 4.67 | 4.3458 | > 5 | 4.274 | 5.025 | 4.37 | 4.5464 | 5.11 | 4.526 | 4.81 |  | 16
| 1 | 4 | 4.47 | 4.77 | 4.664 | 4.4895 | 4.418 | 4.6157 | 4.57 | 4.58 | 4.58 | 4.71 | 4.3652 | > 5 | 4.273 | 4.673 | 4.37 | 4.5075 | 4.54 | 4.461 | 4.86 |  | 4.5728
| 1 | 5 | 4.47 | 4.8 | 4.68 | 4.5135 | 4.394 | 4.7696 | 4.65 | 4.48 | 4.63 | 4.53 | 4.3364 | > 5 | 4.278 | 4.953 | 4.37 | 4.3948 | 4.91 | 4.578 | 4.83 |  | 11
| 1 | 6 | 4.48 | 4.79 | 4.672 | 4.4656 | 4.394 | 4.5214 | 4.65 | 4.41 | 4.59 | 4.59 | 4.3387 | > 5 | 4.274 | 4.809 | 4.36 | 4.4541 | 4.91 | 4.494 | 4.87
| Mean |  | 4.475 | 4.78 | 4.6715 | 4.5268 | 4.3753 | 4.6538 | 4.64 | 4.5283 | 4.585 | 4.6117 | 4.3442 |  | 4.263 | 4.8197 | 4.37 | 4.4655 | 4.8433 | 4.5268 | 4.83
| Median |  | 4.475 | 4.78 | 4.6715 | 4.5215 | 4.382 | 4.6578 | 4.65 | 4.555 | 4.585 | 4.62 | 4.3422 |  | 4.2735 | 4.797 | 4.37 | 4.4808 | 4.9 | 4.5215 | 4.835
| Std Dev. |  | 0.0054772 | 0.02 | 0.0051284 | 0.047571 | 0.036015 | 0.08398 | 0.034641 | 0.06969 | 0.02881 | 0.077567 | 0.015049 |  | 0.018612 | 0.14441 | 0.0063246 | 0.066432 | 0.19735 | 0.047981 | 0.036332
| Rel.Std.Dev. |  | 0.1224% | 0.4184% | 0.1098% | 1.051% | 0.8231% | 1.805% | 0.7466% | 1.539% | 0.6283% | 1.682% | 0.3464% |  | 0.4366% | 2.996% | 0.1447% | 1.488% | 4.075% | 1.06% | 0.7522%
| PDM3 |  | -2.138% | 4.532% | 2.159% | -1.005% | -4.318% | 1.773% | 1.47% | -0.9716% | 0.2676% | 0.8507% | -4.998% |  | -6.774% | 5.399% | -4.434% | -2.346% | 5.917% | -1.004% | 5.625%
| Z-Score (Absolute) |  | 0.52 | 1.42 | 0.73 | 0.19 | 1.16 | 0.62 | 0.53 | 0.18 | 0.18 | 0.35 | 1.35 | NA | 1.87 | 1.67 | 1.19 | 0.58 | 1.82 | 0.19 | 1.74
| Analytical results for Dy in OREAS 555b (Indicative Value 6.71 ppm) | [21 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 6.471 | 7 | [18 empty] | 1
| 1 | 2 | 6.17 | 6.94 | [18 empty] | 10
| 1 | 3 | 6.345 | 7.47 | [18 empty] | 16
| 1 | 4 | 6.367 | 7.21 | [18 empty] | 6.7094
| 1 | 5 | 6.382 | 6.79 | [18 empty] | 16
| 1 | 6 | 6.388 | 6.98
| Mean |  | 6.3538 | 7.065
| Median |  | 6.3745 | 6.99
| Std Dev. |  | 0.099734 | 0.2399
| Rel.Std.Dev. |  | 1.57% | 3.396%
| PDM3 |  | -5.3% | 5.3%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Er in OREAS 555b (Indicative Value 3.87 ppm) | [21 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 3.697 | 4.06 | [18 empty] | 1
| 1 | 2 | 3.437 | 4.05 | [18 empty] | 11
| 1 | 3 | 3.624 | 4.22 | [18 empty] | 16
| 1 | 4 | 3.656 | 4.25 | [18 empty] | 3.8665
| 1 | 5 | 3.676 | 4.07 | [18 empty] | 17
| 1 | 6 | 6.682 | 4.04
| Mean |  | 4.1287 | 4.115
| Median |  | 3.666 | 4.065
| Std Dev. |  | 1.2544 | 0.093968
| Rel.Std.Dev. |  | 30.38% | 2.284%
| PDM3 |  | 6.78% | 6.427%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 555b (Indicative Value 0.89 ppm) | [21 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.877 | 0.88 | [18 empty] | 1
| 1 | 2 | 0.832 | 0.97 | [18 empty] | 12
| 1 | 3 | 0.853 | 0.97 | [18 empty] | 16
| 1 | 4 | 0.855 | 0.96 | [18 empty] | 0.89117
| 1 | 5 | 0.859 | 0.88 | [18 empty] | 18
| 1 | 6 | 0.858 | 0.9
| Mean |  | 0.85567 | 0.92667
| Median |  | 0.8565 | 0.93
| Std Dev. |  | 0.014418 | 0.044572
| Rel.Std.Dev. |  | 1.685% | 4.81%
| PDM3 |  | -3.984% | 3.984%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Fe2O3 in OREAS 555b (Certified Value 5.55 wt.%) | [21 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 4.83 | 5.84 | 5.59 | 5.58 | 5.5 | 5.4 | 5.576 | 5.55 | 5.533 | 6.691 | 5.63 | 5.6562 | 5.2 | 5.4 | 5.79 | 5.78 | 5.217 | 5.05 | 5.6 | 5.619 | 1
| 1 | 2 | 4.75 | 5.82 | 5.56 | 5.56 | 5.5 | 5.41 | 5.59 | 5.6 | 5.662 | 6.72 | 5.63 | 5.6146 | 5.2 | 5.51 | 5.7 | 5.79 | 5.345 | 5.16 | 5.51 | 5.604 | 43
| 1 | 3 | 4.74 | 5.86 | 5.6 | 5.5 | 5.5 | 5.41 | 5.576 | 5.55 | 5.69 | 6.748 | 5.68 | 5.6887 | 5.2 | 5.42 | 5.69 | 5.8 | 5.322 | 5.2 | 5.52 | 5.619 | 16
| 1 | 4 | 4.58 | 5.86 | 5.57 | 5.55 | 5.52 | 5.42 | 5.504 | 5.57 | 5.676 | 6.748 | 5.66 | 5.6727 | 5.2 | 5.35 | 5.68 | 5.79 | 5.319 | 5.33 | 5.48 | 5.59 | 5.5509
| 1 | 5 | 4.83 | 5.87 | 5.58 | 5.56 | 5.49 | 5.43 | 5.576 | 5.64 | 5.547 | 6.762 | 5.64 | 5.6887 | 5.2 | 5.54 | 5.67 | 5.79 | 5.242 | 5.84 | 5.57 | 5.576 | 12
| 1 | 6 | 4.77 | 5.86 | 5.58 | 5.52 | 5.5 | 5.47 | 5.604 | 5.61 | 5.59 | 6.705 | 5.67 | 5.6747 | 5.2 | 5.31 | 5.64 | 5.78 | 5.406 | 5.68 | 5.5 | 5.604
| Mean |  | 4.75 | 5.8517 | 5.58 | 5.545 | 5.5017 | 5.4233 | 5.571 | 5.5867 | 5.6163 | 6.729 | 5.6517 | 5.6659 | 5.2 | 5.4217 | 5.695 | 5.7883 | 5.3085 | 5.3767 | 5.53 | 5.602
| Median |  | 4.76 | 5.86 | 5.58 | 5.555 | 5.5 | 5.415 | 5.576 | 5.585 | 5.626 | 6.734 | 5.65 | 5.6737 | 5.2 | 5.41 | 5.685 | 5.79 | 5.3205 | 5.265 | 5.515 | 5.604
| Std Dev. |  | 0.091869 | 0.018348 | 0.014142 | 0.029496 | 0.0098319 | 0.025033 | 0.034681 | 0.036148 | 0.068582 | 0.027971 | 0.02137 | 0.027881 | 0 | 0.089312 | 0.050892 | 0.0075277 | 0.069174 | 0.31424 | 0.045607 | 0.016769
| Rel.Std.Dev. |  | 1.934% | 0.3136% | 0.2534% | 0.5319% | 0.1787% | 0.4616% | 0.6225% | 0.647% | 1.221% | 0.4157% | 0.3781% | 0.4921% | 0% | 1.647% | 0.8936% | 0.13% | 1.303% | 5.845% | 0.8247% | 0.2993%
| PDM3 |  | -14.43% | 5.419% | 0.5249% | -0.1056% | -0.8863% | -2.297% | 0.3628% | 0.645% | 1.179% | 21.22% | 1.816% | 2.073% | -6.321% | -2.327% | 2.597% | 4.278% | -4.366% | -3.138% | -0.3758% | 0.9213%
| Z-Score (Absolute) |  | 5.31 | 1.77 | 0.03 | 0.2 | 0.47 | 0.98 | 0.03 | 0.07 | 0.26 | 7.41 | 0.49 | 0.58 | 2.41 | 0.99 | 0.77 | 1.37 | 1.72 | 1.28 | 0.29 | 0.17
| Analytical results for Ga in OREAS 555b (Indicative Value 28.6 ppm) | [21 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 1
1
| 1 | 1 | < 100 | 28.3 | [18 empty] | 1
| 1 | 2 | < 100 | 29 | [18 empty] | 13
| 1 | 3 | < 100 | 28.7 | [18 empty] | 16
| 1 | 4 | < 100 | 28.2 | [18 empty] | 28.5833
| 1 | 5 | < 100 | 28.5 | [18 empty] | 19
| 1 | 6 | < 100 | 28.8
| Mean |  |  | 28.5833
| Median |  |  | 28.6
| Std Dev. |  |  | 0.30605
| Rel.Std.Dev. |  |  | 1.071%
| PDM3 |  |  | 1.11e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Gd in OREAS 555b (Indicative Value 7.11 ppm) | [21 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 6.755 | 7.15 | [18 empty] | 1
| 1 | 2 | 6.519 | 7.58 | [18 empty] | 14
| 1 | 3 | 6.647 | 7.92 | [18 empty] | 16
| 1 | 4 | 6.724 | 7.68 | [18 empty] | 7.1144
| 1 | 5 | 6.735 | 7.27 | [18 empty] | 20
| 1 | 6 | 6.743 | 7.65
| Mean |  | 6.6872 | 7.5417
| Median |  | 6.7295 | 7.615
| Std Dev. |  | 0.090837 | 0.28379
| Rel.Std.Dev. |  | 1.358% | 3.763%
| PDM3 |  | -6.005% | 6.005%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ge in OREAS 555b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Ge | Round | Replicate | Lab | [19 empty] | 1
| No | No | 01 | [19 empty] | ppm
|  |  | PPP*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 15
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 21
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 555b (Indicative Value 253 ppm) | [21 empty] | Indicative
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 06 | 08 | 12 | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | 700 | 8.61 | [15 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | 700 | 9.08 | [15 empty] | 1
| 1 | 3 | < 100 | < 100 | 100 | 700 | 8.84 | [15 empty] | 16
| 1 | 4 | < 100 | < 100 | 100 | 600 | 9.08 | [15 empty] | 252.9417
| 1 | 5 | < 100 | < 100 | < 100 | 700 | 8.61 | [15 empty] | 7
| 1 | 6 | < 100 | < 100 | < 100 | 700 | 8.73
| Mean |  |  |  | 100 | 683.3333 | 8.825
| Median |  |  |  | 100 | 700 | 8.785
| Std Dev. |  |  |  | 0 | 40.8248 | 0.21529
| Rel.Std.Dev. |  |  |  | 0% | 5.974% | 2.44%
| PDM3 |  |  |  | -60.47% | 170.2% | -96.51%
| Z-Score (Absolute) |  | 0 | 0 | 0.67 | 25.62 | 1.67
| Analytical results for Hg in OREAS 555b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | [19 empty] | 1
| No | No | 01 | [19 empty] | ppm
|  |  | PPP*XRF | [19 empty] | 0
0
| 1 | 1 | < 100 | [19 empty] | 1
| 1 | 2 | < 100 | [19 empty] | 20
| 1 | 3 | < 100 | [19 empty] | 16
| 1 | 4 | < 100 | [19 empty] | < 100
| 1 | 5 | < 100 | [19 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 555b (Indicative Value 1.39 ppm) | [21 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 1.375 | 1.4 | [18 empty] | 1
| 1 | 2 | 1.323 | 1.43 | [18 empty] | 3
| 1 | 3 | 1.354 | 1.49 | [18 empty] | 16
| 1 | 4 | 1.362 | 1.4 | [18 empty] | 1.387
| 1 | 5 | 1.367 | 1.37 | [18 empty] | 9
| 1 | 6 | 1.363 | 1.41
| Mean |  | 1.3573 | 1.4167
| Median |  | 1.3625 | 1.405
| Std Dev. |  | 0.018162 | 0.040825
| Rel.Std.Dev. |  | 1.338% | 2.882%
| PDM3 |  | -2.139% | 2.139%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 555b (Indicative Value 2.78 ppm) | [21 empty] | Indicative
In | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | < 100 | 2.7 | [18 empty] | 1
| 1 | 2 | < 100 | 2.9 | [18 empty] | 4
| 1 | 3 | < 100 | 2.8 | [18 empty] | 16
| 1 | 4 | < 100 | 2.8 | [18 empty] | 2.7833
| 1 | 5 | < 100 | 2.7 | [18 empty] | 10
| 1 | 6 | < 100 | 2.8
| Mean |  |  | 2.7833
| Median |  |  | 2.8
| Std Dev. |  |  | 0.075277
| Rel.Std.Dev. |  |  | 2.705%
| PDM3 |  |  | 1.11e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for K2O in OREAS 555b (Certified Value 3.08 wt.%) | [21 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 3.59 | 3.08 | 3.1 | 3.07 | 3.12 | 2.98 | 3.18 | 3.09 | 3.06 | 3.09 | 3.0975 | 2.8 | 2.87 | 3.1 | 3.56 | 3.02 | 3.09 | 3.119 | 3.26 |  | 1
| 1 | 2 | 3.51 | 3.08 | 3.1 | 3.03 | 3.09 | 2.99 | 3.14 | 3.09 | 3.13 | 3.11 | 3.0901 | 2.8 | 2.93 | 3.12 | 3.57 | 3.031 | 3.16 | 3.098 | 3.27 |  | 31
| 1 | 3 | 3.58 | 3.09 | 3.1 | 3.02 | 3.08 | 2.97 | 3.15 | 3.08 | 3.12 | 3.13 | 3.1298 | 2.8 | 2.9 | 3.02 | 3.55 | 3.035 | 2.99 | 3.092 | 3.22 |  | 16
| 1 | 4 | 3.58 | 3.08 | 3.08 | 2.97 | 3.13 | 2.96 | 3.12 | 3.09 | 3.14 | 3.1 | 3.1149 | 2.8 | 2.85 | 3.11 | 3.55 | 3.02 | 3.04 | 3.089 | 3.28 |  | 3.079
| 1 | 5 | 3.46 | 3.09 | 3.1 | 3 | 3.09 | 2.98 | 3.19 | 3.11 | 3.07 | 3.11 | 3.106 | 2.8 | 2.95 | 3.02 | 3.55 | 3.02 | 3.18 | 3.109 | 3.28 |  | 13
| 1 | 6 | 3.6 | 3.09 | 3.1 | 3 | 3.12 | 3 | 3.15 | 3.09 | 3.1 | 3.09 | 3.0891 | 2.8 | 2.76 | 3 | 3.55 | 3.041 | 3 | 3.085 | 3.3
| Mean |  | 3.5533 | 3.085 | 3.0967 | 3.015 | 3.105 | 2.98 | 3.155 | 3.0917 | 3.1033 | 3.105 | 3.1045 | 2.8 | 2.8767 | 3.0617 | 3.555 | 3.0278 | 3.0767 | 3.0987 | 3.2683
| Median |  | 3.58 | 3.085 | 3.1 | 3.01 | 3.105 | 2.98 | 3.15 | 3.09 | 3.11 | 3.105 | 3.1017 | 2.8 | 2.885 | 3.06 | 3.55 | 3.0255 | 3.065 | 3.095 | 3.275
| Std Dev. |  | 0.055737 | 0.0054772 | 0.008165 | 0.033912 | 0.020736 | 0.014142 | 0.025884 | 0.0098319 | 0.03266 | 0.015166 | 0.015761 | 4.8648e-16 | 0.06802 | 0.053821 | 0.0083666 | 0.0091524 | 0.080664 | 0.013003 | 0.027142
| Rel.Std.Dev. |  | 1.569% | 0.1775% | 0.2637% | 1.125% | 0.6678% | 0.4746% | 0.8204% | 0.318% | 1.052% | 0.4884% | 0.5077% | 1.737e-14% | 2.365% | 1.758% | 0.2353% | 0.3023% | 2.622% | 0.4196% | 0.8304%
| PDM3 |  | 15.41% | 0.1948% | 0.5737% | -2.079% | 0.8443% | -3.215% | 2.468% | 0.4113% | 0.7902% | 0.8443% | 0.8296% | -9.061% | -6.571% | -0.5631% | 15.46% | -1.662% | -0.07589% | 0.6386% | 6.149%
| Z-Score (Absolute) |  | 8.8 | 0.22 | 0 | 1.57 | 0.16 | 2.25 | 1.12 | 0.1 | 0.13 | 0.16 | 0.15 | 5.72 | 4.24 | 0.67 | 8.83 | 1.33 | 0.39 | 0.04 | 3.31
| Analytical results for La in OREAS 555b (Indicative Value 78 ppm) | [21 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 12 | 13 | 25 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 90 | < 90 | 77.286 | 76.4 | [16 empty] | 1
| 1 | 2 | 170 | 90 | 77.079 | 78.4 | [16 empty] | 5
| 1 | 3 | 90 | 90 | 77.254 | 77.7 | [16 empty] | 16
| 1 | 4 | 90 | 90 | 77.112 | 76.7 | [16 empty] | 77.6183
| 1 | 5 | 90 | < 90 | 77.136 | 74.9 | [16 empty] | 11
| 1 | 6 | 90 | 90 | 77.147 | 77.8
| Mean |  | 103.3333 | 90 | 77.169 | 76.9833
| Median |  | 90 | 90 | 77.1415 | 77.2
| Std Dev. |  | 32.6599 | 0 | 0.082263 | 1.2608
| Rel.Std.Dev. |  | 31.61% | 0% | 0.1066% | 1.638%
| PDM3 |  | 33.13% | 15.95% | -0.5789% | -0.8181%
| Z-Score (Absolute) |  | 4.94 | 1.33 | 0.02 | 0.02
| Analytical results for Lu in OREAS 555b (Indicative Value 0.58 ppm) | [21 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.514 | 0.63 | [18 empty] | 1
| 1 | 2 | 0.503 | 0.65 | [18 empty] | 25
| 1 | 3 | 0.512 | 0.65 | [18 empty] | 16
| 1 | 4 | 0.514 | 0.61 | [18 empty] | 0.57725
| 1 | 5 | 0.516 | 0.62 | [18 empty] | 12
| 1 | 6 | 0.518 | 0.69
| Mean |  | 0.51283 | 0.64167
| Median |  | 0.514 | 0.64
| Std Dev. |  | 0.0052313 | 0.028577
| Rel.Std.Dev. |  | 1.02% | 4.454%
| PDM3 |  | -11.16% | 11.16%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 555b (Certified Value 5.13 wt.%) | [21 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 5.14 | 5.09 | 5.16 | 5.14 | 4.99 | 5.1 | 5.18 | 5.2 | 5.15 | 5.16 | 5.2789 | 4.6 | 5.13 | 5.04 | 4.9 | 5.014 | 4.95 | 5.28 | 5.25 |  | 1
| 1 | 2 | 5.39 | 5.07 | 5.13 | 5.12 | 4.99 | 5.1 | 5.17 | 5.19 | 5.26 | 5.18 | 5.2238 | 4.7 | 5.15 | 5.08 | 5 | 5.0068 | 4.91 | 5.23 | 5.24 |  | 34
| 1 | 3 | 5.43 | 5.08 | 5.16 | 5.09 | 4.97 | 5.09 | 5.16 | 5.17 | 5.27 | 5.21 | 5.2831 | 4.7 | 4.99 | 5.06 | 5 | 5.0037 | 4.77 | 5.23 | 5.24 |  | 16
| 1 | 4 | 5.44 | 5.03 | 5.15 | 5.11 | 5.01 | 5.11 | 5.07 | 5.21 | 5.27 | 5.16 | 5.2779 | 4.7 | 5.08 | 5.05 | 5 | 5.0243 | 5.18 | 5.22 | 5.26 |  | 5.1339
| 1 | 5 | 5.33 | 5.05 | 5.15 | 5.13 | 4.96 | 5.08 | 5.16 | 5.21 | 5.18 | 5.17 | 5.3112 | 4.7 | 5.12 | 5.01 | 5 | 5.012 | 5.5 | 5.24 | 5.27 |  | 14
| 1 | 6 | 5.28 | 5.07 | 5.13 | 5.1 | 5.01 | 5.11 | 5.15 | 5.2 | 5.2 | 5.17 | 5.2269 | 4.7 | 4.95 | 5.02 | 5 | 5.0018 | 5.17 | 5.24 | 5.26
| Mean |  | 5.335 | 5.065 | 5.1467 | 5.115 | 4.9883 | 5.0983 | 5.1483 | 5.1967 | 5.2217 | 5.175 | 5.267 | 4.6833 | 5.07 | 5.0433 | 4.9833 | 5.0105 | 5.08 | 5.24 | 5.2533
| Median |  | 5.36 | 5.07 | 5.15 | 5.115 | 4.99 | 5.1 | 5.16 | 5.2 | 5.23 | 5.17 | 5.2784 | 4.7 | 5.1 | 5.045 | 5 | 5.0094 | 5.06 | 5.235 | 5.255
| Std Dev. |  | 0.11327 | 0.021679 | 0.013663 | 0.018708 | 0.020412 | 0.01169 | 0.039707 | 0.015055 | 0.051929 | 0.018708 | 0.034473 | 0.040825 | 0.081731 | 0.02582 | 0.040825 | 0.0082608 | 0.25938 | 0.020976 | 0.012111
| Rel.Std.Dev. |  | 2.123% | 0.428% | 0.2655% | 0.3658% | 0.4092% | 0.2293% | 0.7713% | 0.2897% | 0.9945% | 0.3615% | 0.6545% | 0.8717% | 1.612% | 0.512% | 0.8192% | 0.1649% | 5.106% | 0.4003% | 0.2305%
| PDM3 |  | 3.917% | -1.342% | 0.249% | -0.3678% | -2.835% | -0.6925% | 0.2814% | 1.223% | 1.71% | 0.8009% | 2.592% | -8.776% | -1.244% | -1.764% | -2.932% | -2.404% | -1.05% | 2.067% | 2.327%
| Z-Score (Absolute) |  | 1.82 | 0.41 | 0.26 | 0 | 1.05 | 0.14 | 0.28 | 0.67 | 0.88 | 0.5 | 1.25 | 3.56 | 0.37 | 0.59 | 1.09 | 0.86 | 0.29 | 1.03 | 1.14
| Analytical results for MnO in OREAS 555b (Certified Value 0.117 wt.%) | [21 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.13 | 0.12 | 0.12 | 0.124 | 0.11 | 0.113 | 0.116 | 0.12 | 0.116 | 0.13429 | 0.12 | 0.12213 | 0.1 | 0.11 | 0.035947 | 0.116 | 0.1126 | 0.118 | 0.119 | 0.129 | 1
| 1 | 2 | 0.11 | 0.12 | 0.12 | 0.121 | 0.12 | 0.112 | 0.116 | 0.12 | 0.116 | 0.13041 | 0.11 | 0.12159 | 0.1 | 0.11 | 0.04044 | 0.116 | 0.11913 | 0.114 | 0.114 | 0.116 | 35
| 1 | 3 | 0.1 | 0.12 | 0.12 | 0.119 | 0.13 | 0.112 | 0.116 | 0.11 | 0.116 | 0.13558 | 0.11 | 0.12051 | 0.1 | 0.12 | 0.04044 | 0.116 | 0.11342 | 0.118 | 0.117 | 0.116 | 16
| 1 | 4 | 0.11 | 0.12 | 0.12 | 0.123 | 0.11 | 0.117 | 0.116 | 0.12 | 0.116 | 0.13429 | 0.11 | 0.12159 | 0.1 | 0.11 | 0.044933 | 0.116 | 0.13463 | 0.116 | 0.114 | 0.116 | 0.11679
| 1 | 5 | 0.11 | 0.12 | 0.12 | 0.121 | 0.12 | 0.113 | 0.116 | 0.12 | 0.116 | 0.133 | 0.11 | 0.12804 | 0.1 | 0.12 | 0.04044 | 0.116 | 0.10689 | 0.129 | 0.117 | 0.116 | 15
| 1 | 6 | 0.12 | 0.12 | 0.12 | 0.12 | 0.11 | 0.114 | 0.116 | 0.12 | 0.116 | 0.13558 | 0.11 | 0.12266 | 0.1 | 0.11 | 0.04044 | 0.116 | 0.11505 | 0.121 | 0.114 | 0.116
| Mean |  | 0.11333 | 0.12 | 0.12 | 0.12133 | 0.11667 | 0.1135 | 0.116 | 0.11833 | 0.116 | 0.13386 | 0.11167 | 0.12275 | 0.1 | 0.11333 | 0.04044 | 0.116 | 0.11695 | 0.11933 | 0.11583 | 0.11817
| Median |  | 0.11 | 0.12 | 0.12 | 0.121 | 0.115 | 0.113 | 0.116 | 0.12 | 0.116 | 0.13429 | 0.11 | 0.12186 | 0.1 | 0.11 | 0.04044 | 0.116 | 0.11424 | 0.118 | 0.1155 | 0.116
| Std Dev. |  | 0.010328 | 0 | 0 | 0.0018619 | 0.008165 | 0.0018708 | 0 | 0.0040825 | 0 | 0.0019457 | 0.0040825 | 0.0026882 | 1.5202e-17 | 0.005164 | 0.0028418 | 0 | 0.009525 | 0.0052789 | 0.002137 | 0.0053072
| Rel.Std.Dev. |  | 9.113% | 0% | 0% | 1.535% | 6.999% | 1.648% | 0% | 3.45% | 0% | 1.454% | 3.656% | 2.19% | 1.52e-14% | 4.556% | 7.027% | 0% | 8.144% | 4.424% | 1.845% | 4.491%
| PDM3 |  | -2.962% | 2.746% | 2.746% | 3.888% | -0.1079% | -2.819% | -0.6787% | 1.319% | -0.6787% | 14.61% | -4.389% | 5.104% | -14.38% | -2.962% | -65.37% | -0.6787% | 0.139% | 2.175% | -0.8214% | 1.176%
| Z-Score (Absolute) |  | 0.71 | 0.71 | 0.71 | 0.99 | 0 | 0.67 | 0.1 | 0.35 | 0.1 | 3.66 | 1.06 | 1.3 | NA | 0.71 | 16.23 | 0.1 | 0.06 | 0.57 | 0.18 | 0.36
Results from laboratory 19 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Mo in OREAS 555b (Indicative Value < 50 ppm) | [21 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 08 | 12 | 25 | [16 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [16 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | 16.4 | [16 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | 17.5 | [16 empty] | 7
| 1 | 3 | < 100 | < 50 | < 50 | 16.9 | [16 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | 17 | [16 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | 17 | [16 empty] | 13
| 1 | 6 | < 100 | < 50 | < 50 | 16.6
| Mean |  |  |  |  | 16.9
| Median |  |  |  |  | 16.95
| Std Dev. |  |  |  |  | 0.37947
| Rel.Std.Dev. |  |  |  |  | 2.245%
PDM3
| Z-Score (Absolute) |  | 4.16 | 0 | 0 | 1.35
| Analytical results for Na2O in OREAS 555b (Indicative Value 0.103 wt.%) | [21 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 17 | 19 | 21 | 22 | 25 | 28 | [10 empty] | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 3
3
| 1 | 1 | < 0.01 | 0.13 | 0.04 | < 0.01 | 0.15 | < 0.1 | 0.07 | 0.08 | 0.6 | 0.05 | [10 empty] | 1
| 1 | 2 | < 0.01 | 0.12 | 0.05 | < 0.01 | 0.15 | < 0.1 | 0.07 | 0.07 | 0.6 | 0.05 | [10 empty] | 8
| 1 | 3 | < 0.01 | 0.11 | 0.05 | < 0.01 | 0.16 | < 0.1 | 0.07 | 0.08 | 0.39 | 0.05 | [10 empty] | 16
| 1 | 4 | < 0.01 | 0.13 | 0.07 | < 0.01 | 0.13 | < 0.1 | 0.07 | 0.08 | 0.35 | 0.05 | [10 empty] | 0.10333
| 1 | 5 | < 0.01 | 0.13 | 0.06 | < 0.01 | 0.16 | 0.2 | 0.07 | 0.06 | 0.17 | 0.05 | [10 empty] | 14
| 1 | 6 | < 0.01 | 0.12 | 0.06 | < 0.01 | 0.15 | < 0.1 | 0.06 | 0.08 | 0.43 | 0.06
| Mean |  |  | 0.12333 | 0.055 |  | 0.15 | 0.2 | 0.068333 | 0.075 | 0.42333 | 0.051667
| Median |  |  | 0.125 | 0.055 |  | 0.15 | 0.2 | 0.07 | 0.08 | 0.41 | 0.05
| Std Dev. |  |  | 0.008165 | 0.010488 |  | 0.010954 |  | 0.0040825 | 0.0083666 | 0.16318 | 0.0040825
| Rel.Std.Dev. |  |  | 6.62% | 19.07% |  | 7.303% |  | 5.974% | 11.16% | 38.55% | 7.902%
| PDM3 |  |  | 19.35% | -46.77% |  | 45.16% | 93.55% | -33.87% | -27.42% | 309.7% | -50%
| Z-Score (Absolute) |  | 1.25 | 0.97 | 0.31 | 1.25 | 1.47 | 0.06 | 0.06 | 0.06 | 6.62 | 0.38
| Analytical results for Nb in OREAS 555b (Indicative Value 40 ppm) | [21 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 12 | 21 | 25 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*OES | BF*XRF | [16 empty] | 1
1
| 1 | 1 | < 50 | 50 | 41 | 25.7 | [16 empty] | 1
| 1 | 2 | < 50 | 50 | 44 | 25.7 | [16 empty] | 9
| 1 | 3 | < 50 | 60 | 42 | 25.7 | [16 empty] | 16
| 1 | 4 | < 50 | 70 | 41 | 26.3 | [16 empty] | 40.0056
| 1 | 5 | < 50 | 60 | 40 | 25.8 | [16 empty] | 15
| 1 | 6 | < 50 | < 50 | 42 | 25.9
| Mean |  |  | 58 | 41.6667 | 25.85
| Median |  |  | 60 | 41.5 | 25.75
| Std Dev. |  |  | 8.3666 | 1.3663 | 0.23452
| Rel.Std.Dev. |  |  | 14.43% | 3.279% | 0.9072%
| PDM3 |  |  | 44.98% | 4.152% | -35.38%
| Z-Score (Absolute) |  | 0.71 | 1.52 | 0.64 | 0.64
| Analytical results for Nd in OREAS 555b (Indicative Value 63 ppm) | [21 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 0
0
| 1 | 1 | 62.904 | 60.7 | [18 empty] | 1
| 1 | 2 | 62.439 | 64.6 | [18 empty] | 10
| 1 | 3 | 62.864 | 63.3 | [18 empty] | 16
| 1 | 4 | 62.763 | 61.1 | [18 empty] | 62.5805
| 1 | 5 | 62.648 | 60.9 | [18 empty] | 16
| 1 | 6 | 62.748 | 64
| Mean |  | 62.7277 | 62.4333
| Median |  | 62.7555 | 62.2
| Std Dev. |  | 0.16794 | 1.734
| Rel.Std.Dev. |  | 0.2677% | 2.777%
| PDM3 |  | 0.2352% | -0.2352%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 555b (Certified Value 233 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 08 | 12 | 17 | 22 | 24 | 25 | 27 | 28 | [8 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 200 | 200 | 212 | 160 | 280 | 230 | 290 | 240 | 220.4322 | 230 | 250 | < 100 | [8 empty] | 1
| 1 | 2 | 200 | 100 | 212 | 160 | 220 | 230 | 190 | 240 | 227.112 | 237 | 240 | 100 | [8 empty] | 40
| 1 | 3 | 200 | 100 | 244 | 160 | 230 | 260 | 220 | 240 | 233.7917 | 246 | 240 | 100 | [8 empty] | 16
| 1 | 4 | 200 | 100 | 196 | 160 | 280 | 240 | 310 | 240 | 221.9166 | 235 | 240 | < 100 | [8 empty] | 232.5519
| 1 | 5 | 200 | 200 | 204 | 160 | 250 | 240 | 160 | 240 | 207.0727 | 235 | 290 | 100 | [8 empty] | 16
| 1 | 6 | 200 | 100 | 291 | 160 | 250 | 220 | 160 | 240 | 227.112 | 233 | 240 | 100
| Mean |  | 200 | 133.3333 | 226.5 | 160 | 251.6667 | 236.6667 | 221.6667 | 240 | 222.9062 | 236 | 250 | 100
| Median |  | 200 | 100 | 212 | 160 | 250 | 235 | 205 | 240 | 224.5143 | 235 | 240 | 100
| Std Dev. |  | 0 | 51.6398 | 35.5626 | 0 | 24.8328 | 13.6626 | 64.9359 | 0 | 9.0738 | 5.4406 | 20 | 0
| Rel.Std.Dev. |  | 0% | 38.73% | 15.7% | 0% | 9.867% | 5.773% | 29.29% | 0% | 4.071% | 2.305% | 8% | 0%
| PDM3 |  | -14% | -42.67% | -2.602% | -31.2% | 8.22% | 1.769% | -4.681% | 3.203% | -4.148% | 1.483% | 7.503% | -57%
| Z-Score (Absolute) |  | NA | NA | 0.48 | 4.12 | 0.83 | 0.05 | 0.74 | 0 | 0.67 | 0.01 | 0.75 | NA
Results from laboratories 1, 3 and 28 were removed due to their 100 ppm reading resolution.
| Analytical results for P2O5 in OREAS 555b (Certified Value 0.083 wt.%) | [21 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.07 | 0.08 | 0.09 | 0.09 | 0.08 | 0.078 | 0.08 | 0.09 | 0.08 | 0.074 | 0.078551 | < 0.1 | 0.1 | 0.08 | 0.1 | 0.085 | 0.11 | 0.085 | 0.092 |  | 1
| 1 | 2 | 0.06 | 0.08 | 0.081 | 0.09 | 0.08 | 0.078 | 0.08 | 0.1 | 0.08 | 0.075 | 0.076481 | < 0.1 | 0.1 | 0.08 | 0.1 | 0.089 | 0.12 | 0.084 | 0.092 |  | 41
| 1 | 3 | 0.09 | 0.08 | 0.086 | 0.09 | 0.08 | 0.078 | 0.08 | 0.1 | 0.09 | 0.075 | 0.079586 | < 0.1 | 0.11 | 0.08 | 0.1 | 0.086 | 0.1 | 0.084 | 0.092 |  | 16
| 1 | 4 | 0.1 | 0.08 | 0.086 | 0.09 | 0.08 | 0.078 | 0.08 | 0.09 | 0.08 | 0.073 | 0.077516 | < 0.1 | 0.1 | 0.08 | 0.11 | 0.087 | 0.15 | 0.083 | 0.069 |  | 0.083499
| 1 | 5 | 0.09 | 0.08 | 0.086 | 0.09 | 0.08 | 0.078 | 0.08 | 0.09 | 0.08 | 0.073 | 0.079586 | < 0.1 | 0.1 | 0.08 | 0.11 | 0.088 | 0.08 | 0.085 | 0.092 |  | 17
| 1 | 6 | 0.09 | 0.08 | 0.086 | 0.09 | 0.08 | 0.079 | 0.08 | 0.09 | 0.08 | 0.073 | 0.075445 | < 0.1 | 0.1 | 0.08 | 0.11 | 0.086 | 0.09 | 0.085 | 0.092
| Mean |  | 0.083333 | 0.08 | 0.085833 | 0.09 | 0.08 | 0.078167 | 0.08 | 0.093333 | 0.081667 | 0.073833 | 0.077861 |  | 0.10167 | 0.08 | 0.105 | 0.086833 | 0.10833 | 0.084333 | 0.088167
| Median |  | 0.09 | 0.08 | 0.086 | 0.09 | 0.08 | 0.078 | 0.08 | 0.09 | 0.08 | 0.0735 | 0.078033 |  | 0.1 | 0.08 | 0.105 | 0.0865 | 0.105 | 0.0845 | 0.092
| Std Dev. |  | 0.015055 | 0 | 0.0028577 | 1.5202e-17 | 0 | 0.00040825 | 0 | 0.005164 | 0.0040825 | 0.00098319 | 0.0016904 |  | 0.0040825 | 0 | 0.0054772 | 0.001472 | 0.024833 | 0.0008165 | 0.0093897
| Rel.Std.Dev. |  | 18.07% | 0% | 3.329% | 1.689e-14% | 0% | 0.5223% | 0% | 5.533% | 4.999% | 1.332% | 2.171% |  | 4.016% | 0% | 5.216% | 1.695% | 22.92% | 0.9682% | 10.65%
| PDM3 |  | -0.1989% | -4.191% | 2.795% | 7.785% | -4.191% | -6.387% | -4.191% | 11.78% | -2.195% | -11.58% | -6.753% |  | 21.76% | -4.191% | 25.75% | 3.993% | 29.74% | 0.9987% | 5.59%
| Z-Score (Absolute) |  | 0.13 | 0.39 | 0.52 | 1.17 | 0.39 | 0.67 | 0.39 | 1.69 | 0.13 | 1.35 | 0.72 | 5.06 | NA | 0.39 | 3.5 | 0.67 | 4.02 | 0.29 | 0.83
Results from laboratory 21 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 555b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 12 | 13 | 17 | 22 | 25 | 28 | [9 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 90 | < 35 | 50 | < 50 | < 100 | < 50 | 370 | 9.3 | < 100 | [9 empty] | 1
| 1 | 2 | < 100 | < 100 | < 90 | < 35 | < 50 | 80 | < 100 | < 50 | 370 | 9.8 | < 100 | [9 empty] | 11
| 1 | 3 | < 100 | < 100 | < 90 | < 35 | < 50 | 90 | < 100 | < 50 | 370 | 10.4 | < 100 | [9 empty] | 16
| 1 | 4 | < 100 | < 100 | < 90 | < 35 | < 50 | 90 | < 100 | < 50 | 370 | 10.2 | < 100 | [9 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 90 | < 35 | 50 | < 50 | < 100 | < 50 | 370 | 11.9 | < 100 | [9 empty] | 17
| 1 | 6 | < 100 | < 100 | < 90 | < 35 | < 50 | 50 | < 100 | < 50 | 370 | 13.1 | < 100
| Mean | [5 empty] | 50 | 77.5 |  |  | 370 | 10.7833
| Median | [5 empty] | 50 | 85 |  |  | 370 | 10.3
| Std Dev. | [5 empty] | 0 | 18.9297 |  |  | 0 | 1.4331
| Rel.Std.Dev. | [5 empty] | 0% | 24.43% |  |  | 0% | 13.29%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 0.24 | 2.19 | 1.12 | 0.67 | 0 | 1.69 | 21.67 | 2.64 | 0
| Analytical results for Pr in OREAS 555b (Indicative Value 17.4 ppm) | [21 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 1
1
| 1 | 1 | 17.056 | 17.5 | [18 empty] | 1
| 1 | 2 | 17.035 | 18 | [18 empty] | 12
| 1 | 3 | 17.124 | 17.8 | [18 empty] | 16
| 1 | 4 | 17.105 | 17.4 | [18 empty] | 17.3803
| 1 | 5 | 17.117 | 17.7 | [18 empty] | 18
| 1 | 6 | 17.126 | 17.6
| Mean |  | 17.0938 | 17.6667
| Median |  | 17.111 | 17.65
| Std Dev. |  | 0.038727 | 0.21602
| Rel.Std.Dev. |  | 0.2266% | 1.223%
| PDM3 |  | -1.648% | 1.648%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 555b (Indicative Value 103 ppm) | [21 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 08 | 12 | 25 | [16 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 100 | 110 | < 50 | 97.7 | [16 empty] | 1
| 1 | 2 | 100 | 110 | < 50 | 99.1 | [16 empty] | 13
| 1 | 3 | 100 | 110 | < 50 | 102 | [16 empty] | 16
| 1 | 4 | 100 | 110 | < 50 | 98 | [16 empty] | 103.0667
| 1 | 5 | 100 | 110 | < 50 | 98.4 | [16 empty] | 19
| 1 | 6 | 100 | 110 | < 50 | 100
| Mean |  | 100 | 110 |  | 99.2
| Median |  | 100 | 110 |  | 98.75
| Std Dev. |  | 0 | 0 |  | 1.6012
| Rel.Std.Dev. |  | 0% | 0% |  | 1.614%
| PDM3 |  | -2.975% | 6.727% |  | -3.752%
| Z-Score (Absolute) |  | 0.05 | 1.3 | 9.32 | 0.05
| Analytical results for Re in OREAS 555b (Indicative Value 18.3 ppm) | [21 empty] | Indicative
Re | Round | Replicate | Lab | [19 empty] | 1
| No | No | 25 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 1
1
| 1 | 1 | 21 | [19 empty] | 1
| 1 | 2 | 13.4 | [19 empty] | 14
| 1 | 3 | 15 | [19 empty] | 16
| 1 | 4 | 22.6 | [19 empty] | 18.2833
| 1 | 5 | 18 | [19 empty] | 20
| 1 | 6 | 19.7
| Mean |  | 18.2833
| Median |  | 18.85
| Std Dev. |  | 3.542
| Rel.Std.Dev. |  | 19.37%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 555b (Indicative Value < 50 ppm) | [21 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 08 | 12 | 24 | 25 | [15 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 42 | 4 | [15 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 42 | 4 | [15 empty] | 15
| 1 | 3 | < 100 | < 50 | < 50 | < 42 | 4 | [15 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 42 | 5 | [15 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 42 | 3 | [15 empty] | 21
| 1 | 6 | < 100 | < 50 | < 50 | < 42 | 4
| Mean | [5 empty] | 4
| Median | [5 empty] | 4
| Std Dev. | [5 empty] | 0.63246
| Rel.Std.Dev. | [5 empty] | 15.81%
PDM3
| Z-Score (Absolute) |  | 4.1 | 0 | 0 | 0.67 | 3.44
| Analytical results for Sc in OREAS 555b (Indicative Value 13.9 ppm) | [21 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 13 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*MS | BF*XRF | [17 empty] | 1
1
| 1 | 1 | < 100 | 15.554 | 12 | [17 empty] | 1
| 1 | 2 | < 100 | 14.792 | 13 | [17 empty] | 1
| 1 | 3 | < 100 | 15.254 | 13 | [17 empty] | 16
| 1 | 4 | < 100 | 15.241 | 13 | [17 empty] | 13.8577
| 1 | 5 | < 100 | 15.217 | 12 | [17 empty] | 7
| 1 | 6 | < 100 | 15.234 | 12
| Mean |  |  | 15.2153 | 12.5
| Median |  |  | 15.2375 | 12.5
| Std Dev. |  |  | 0.24348 | 0.54772
| Rel.Std.Dev. |  |  | 1.6% | 4.382%
| PDM3 |  |  | 9.797% | -9.797%
| Z-Score (Absolute) |  | 8.64 | 0 | 0.67
| Analytical results for Se in OREAS 555b (Indicative Value 2.67 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | < 100 | 2 | [18 empty] | 1
| 1 | 2 | < 100 | 4 | [18 empty] | 19
| 1 | 3 | < 100 | 4 | [18 empty] | 16
| 1 | 4 | < 100 | 1 | [18 empty] | 2.6667
| 1 | 5 | < 100 | 3 | [18 empty] | 8
| 1 | 6 | < 100 | 2
| Mean |  |  | 2.6667
| Median |  |  | 2.5
| Std Dev. |  |  | 1.2111
| Rel.Std.Dev. |  |  | 45.41%
| PDM3 |  |  | -1.332e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for SiO2 in OREAS 555b (Certified Value 51.77 wt.%) | [21 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 27 | 28 | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 45.79 | 51.78 | 52.3 | 52.12 | 51.55 | 50.81 | 52 | 51.6 | 51.7 | 110.875 | 51.47 | 52.0432 | 44.7 | 53.3 | 51.73 | 51.9 | 50.7929 | 52.89 | 52.64 | 52.04 | 1
| 1 | 2 | 45.51 | 51.69 | 52.16 | 51.56 | 51.35 | 50.86 | 52 | 52 | 52.7 | 112.073 | 51.17 | 51.7786 | 44.8 | 53.24 | 51.96 | 51.9 | 50.8969 | 50.61 | 52.12 | 51.85 | #N/A
| 1 | 3 | 45.4 | 51.53 | 52.31 | 51.24 | 51.32 | 50.78 | 52 | 51.7 | 52.7 | 111.131 | 51.52 | 52.1846 | 44.7 | 53.53 | 50.55 | 52 | 50.7613 | 49.26 | 52.16 | 51.94 | 16
| 1 | 4 | 45.84 | 51.8 | 52.02 | 51.57 | 51.34 | 50.7 | 51.3 | 51.8 | 52.7 | 111.089 | 51.41 | 52.1533 | 44.8 | 52.67 | 50.47 | 51.9 | 50.7514 | 52.19 | 52.17 | 51.86 | 51.7731
| 1 | 5 | 45.74 | 51.82 | 52.26 | 51.82 | 51 | 50.68 | 52 | 51.8 | 51.8 | 111.773 | 51.22 | 52.2866 | 44.8 | 53.43 | 50.61 | 52 | 50.9216 | 55.33 | 52.26 | 51.84 | 18
| 1 | 6 | 45.8 | 51.76 | 52.07 | 51.37 | 51.6 | 51.04 | 51.9 | 51.8 | 52.2 | 111.452 | 51.36 | 51.7937 | 44.9 | 53.23 | 50.11 | 51.9 | 50.7375 | 52.48 | 52.26 | 51.87
| Mean |  | 45.68 | 51.73 | 52.1867 | 51.6133 | 51.36 | 50.8117 | 51.8667 | 51.7833 | 52.3 | 111.3988 | 51.3583 | 52.04 | 44.7833 | 53.2333 | 50.905 | 51.9333 | 50.8103 | 52.1267 | 52.2683 | 51.9
| Median |  | 45.765 | 51.77 | 52.21 | 51.565 | 51.345 | 50.795 | 52 | 51.8 | 52.45 | 111.2915 | 51.385 | 52.0983 | 44.8 | 53.27 | 50.58 | 51.9 | 50.7771 | 52.335 | 52.215 | 51.865
| Std Dev. |  | 0.18055 | 0.1077 | 0.12291 | 0.31709 | 0.21232 | 0.1306 | 0.28048 | 0.13292 | 0.46904 | 0.45595 | 0.13848 | 0.21145 | 0.075277 | 0.29951 | 0.75206 | 0.05164 | 0.079214 | 2.0732 | 0.19062 | 0.077201
| Rel.Std.Dev. |  | 0.3953% | 0.2082% | 0.2355% | 0.6144% | 0.4134% | 0.257% | 0.5408% | 0.2567% | 0.8968% | 0.4093% | 0.2696% | 0.4063% | 0.1681% | 0.5626% | 1.477% | 0.09943% | 0.1559% | 3.977% | 0.3647% | 0.1487%
| PDM3 |  | -11.77% | -0.08326% | 0.7988% | -0.3086% | -0.7979% | -1.857% | 0.1807% | 0.01976% | 1.018% | 115.2% | -0.8011% | 0.5155% | -13.5% | 2.82% | -1.677% | 0.3095% | -1.86% | 0.6829% | 0.9565% | 0.2451%
| Z-Score (Absolute) |  | 9.12 | 0.14 | 0.54 | 0.31 | 0.69 | 1.5 | 0.06 | 0.06 | 0.71 | 88.45 | 0.69 | 0.32 | 10.45 | 2.09 | 1.37 | 0.16 | 1.51 | 0.45 | 0.66 | 0.11
| Analytical results for Sm in OREAS 555b (Indicative Value 9.25 ppm) | [21 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 8.942 | 9.3 | [18 empty] | 1
| 1 | 2 | 8.914 | 10.1 | [18 empty] | 3
| 1 | 3 | 8.936 | 9.6 | [18 empty] | 16
| 1 | 4 | 8.924 | 9.2 | [18 empty] | 9.2526
| 1 | 5 | 8.978 | 8.9 | [18 empty] | 9
| 1 | 6 | 9.837 | 9.4
| Mean |  | 9.0885 | 9.4167
| Median |  | 8.939 | 9.35
| Std Dev. |  | 0.36734 | 0.40702
| Rel.Std.Dev. |  | 4.042% | 4.322%
| PDM3 |  | -1.773% | 1.773%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 555b (Indicative Value 93 ppm) | [21 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 08 | 12 | 13 | 17 | 25 | 28 | [13 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | 60 | 90 | < 100 | 190 | 3.2 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | 50 | 110 | < 100 | 190 | 3.4 | < 100 | [13 empty] | 20
| 1 | 3 | < 100 | 80 | 80 | < 100 | 250 | 3.1 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | 90 | 70 | < 100 | 210 | 3.2 | < 100 | [13 empty] | 93.2292
| 1 | 5 | < 100 | 120 | 100 | < 100 | 200 | 3.5 | < 100 | [13 empty] | 10
| 1 | 6 | < 100 | 90 | 90 | < 100 | 200 | 3.1 | < 100
| Mean |  |  | 81.6667 | 90 |  | 206.6667 | 3.25
| Median |  |  | 85 | 90 |  | 200 | 3.2
| Std Dev. |  |  | 24.8328 | 14.1421 |  | 22.5093 | 0.16432
| Rel.Std.Dev. |  |  | 30.41% | 15.71% |  | 10.89% | 5.056%
| PDM3 |  |  | -12.4% | -3.464% |  | 121.7% | -96.51%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.85 | 0 | 3.34 | 1 | 0
| Analytical results for SO3 in OREAS 555b (Certified Value 11.05 wt.%) | [21 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 08 | 10 | 12 | 17 | 18 | 19 | 22 | 24 | 28 | [7 empty] | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 2
3
| 1 | 1 | 6.76 | 0.16 | 11.57 | 11.56 | 11.512 | 11.4 | 11.337 | > 5 | 11.102 | 10 | 5.32 | 11.0338 | 9.839 | [7 empty] | 1
| 1 | 2 | 6.97 | 0.14 | 11.55 | 11.25 | 11.312 | 11.45 | 11.562 | > 5 | 11.084 | 10.1 | 5.6 | 11.1298 | 10.164 | [7 empty] | #N/A
| 1 | 3 | 6.83 | 0.16 | 11.53 | 11.28 | 11.362 | 11.35 | 11.587 | > 5 | 11.22 | 10 | 5.5 | 11.2599 | 9.939 | [7 empty] | 16
| 1 | 4 | 6.65 | 0.16 | 11.54 | 11.35 | 11.262 | 11.45 | 11.612 | > 5 | 11.144 | 10 | 5.39 | 11.2189 | 9.964 | [7 empty] | 11.0536
| 1 | 5 | 6.96 | 0.16 | 11.6 | 11.36 | 11.512 | 11.5 | 11.337 | > 5 | 11.195 | 10 | 5.74 | 11.0988 | 10.039 | [7 empty] | 19
| 1 | 6 | 6.72 | 0.15 | 11.57 | 11.33 | 11.337 | 11.4 | 11.462 | > 5 | 11.127 | 10 | 5.17 | 11.0868 | 9.914
| Mean |  | 6.815 | 0.155 | 11.56 | 11.355 | 11.3828 | 11.425 | 11.4828 |  | 11.1453 | 10.0167 | 5.4533 | 11.138 | 9.9765
| Median |  | 6.795 | 0.16 | 11.56 | 11.34 | 11.3495 | 11.425 | 11.512 |  | 11.1355 | 10 | 5.445 | 11.1143 | 9.9515
| Std Dev. |  | 0.13004 | 0.0083666 | 0.025298 | 0.10895 | 0.10538 | 0.05244 | 0.12391 |  | 0.052952 | 0.040825 | 0.20393 | 0.085384 | 0.11264
| Rel.Std.Dev. |  | 1.908% | 5.398% | 0.2188% | 0.9595% | 0.9257% | 0.459% | 1.079% |  | 0.4751% | 0.4076% | 3.74% | 0.7666% | 1.129%
| PDM3 |  | -38.35% | -98.6% | 4.581% | 2.727% | 2.979% | 3.36% | 3.883% |  | 0.8301% | -9.381% | -50.66% | 0.7639% | -9.744%
| Z-Score (Absolute) |  | 7.68 | 19.51 | 0.74 | 0.38 | 0.43 | 0.5 | 0.61 | NA | 0.01 | 2 | 10.1 | 0.01 | 2.07
| Analytical results for Sr in OREAS 555b (Certified Value 67 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 06 | 07 | 08 | 10 | 12 | 21 | 22 | 25 | 28 | [10 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 80 | 80 | 68 | < 100 | < 80 | 100 | 75 | < 3400 | 60 | < 100 | [10 empty] | 1
| 1 | 2 | < 80 | < 80 | 67 | < 100 | < 80 | 100 | 79 | < 3400 | 61 | < 100 | [10 empty] | 32
| 1 | 3 | 80 | < 80 | 66 | < 100 | < 80 | 100 | 76 | < 3400 | 60 | < 100 | [10 empty] | 16
| 1 | 4 | < 80 | < 80 | 68 | < 100 | < 80 | 100 | 74 | < 3400 | 63 | < 100 | [10 empty] | 67.1921
| 1 | 5 | < 80 | < 80 | 67 | < 100 | < 80 | 100 | 73 | < 3400 | 65 | < 100 | [10 empty] | 20
| 1 | 6 | < 80 | < 80 | 67 | < 100 | < 80 | 100 | 75 | < 3400 | 63 | < 100
| Mean |  | 80 | 80 | 67.1667 |  |  | 100 | 75.3333 |  | 62
| Median |  | 80 | 80 | 67 |  |  | 100 | 75 |  | 62
| Std Dev. |  |  |  | 0.75277 |  |  | 0 | 2.0656 |  | 2
| Rel.Std.Dev. |  |  |  | 1.121% |  |  | 0% | 2.742% |  | 3.226%
| PDM3 |  | 19.06% | 19.06% | -0.03787% |  |  | 48.83% | 12.12% |  | -7.727%
| Z-Score (Absolute) |  | 0.5 | 0.5 | 0.84 | 0.45 | 1.04 | 3.33 | 1.46 | 123.61 | 0.45 | 0.45
| Analytical results for Ta in OREAS 555b (Indicative Value 1.47 ppm) | [21 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | < 100 | 1.5 | [18 empty] | 1
| 1 | 2 | < 100 | 1.4 | [18 empty] | 5
| 1 | 3 | < 100 | 1.5 | [18 empty] | 16
| 1 | 4 | < 100 | 1.4 | [18 empty] | 1.4667
| 1 | 5 | < 100 | 1.5 | [18 empty] | 11
| 1 | 6 | < 100 | 1.5
| Mean |  |  | 1.4667
| Median |  |  | 1.5
| Std Dev. |  |  | 0.05164
| Rel.Std.Dev. |  |  | 3.521%
| PDM3 |  |  | -2.109e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Tb in OREAS 555b (Indicative Value 1.15 ppm) | [21 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 1.177 | 1.09 | [18 empty] | 1
| 1 | 2 | 1.112 | 1.13 | [18 empty] | 6
| 1 | 3 | 1.127 | 1.12 | [18 empty] | 16
| 1 | 4 | 1.126 | 1.15 | [18 empty] | 1.1494
| 1 | 5 | 1.133 | 1.04 | [18 empty] | 12
| 1 | 6 | 1.448 | 1.14
| Mean |  | 1.1872 | 1.1117
| Median |  | 1.13 | 1.125
| Std Dev. |  | 0.12968 | 0.040702
| Rel.Std.Dev. |  | 10.92% | 3.661%
| PDM3 |  | 3.284% | -3.284%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 555b (Indicative Value 10.9 ppm) | [21 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 1
1
| 1 | 1 | < 100 | 5.5 | [18 empty] | 1
| 1 | 2 | < 100 | 19.7 | [18 empty] | 7
| 1 | 3 | < 100 | 21.7 | [18 empty] | 16
| 1 | 4 | < 100 | 4.9 | [18 empty] | 10.9167
| 1 | 5 | < 100 | 4.9 | [18 empty] | 13
| 1 | 6 | < 100 | 8.8
| Mean |  |  | 10.9167
| Median |  |  | 7.15
| Std Dev. |  |  | 7.7414
| Rel.Std.Dev. |  |  | 70.91%
| PDM3 |  |  | -3.109e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Th in OREAS 555b (Indicative Value 15.7 ppm) | [21 empty] | Indicative
Th | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 1
1
| 1 | 1 | 14.342 | 17 | [18 empty] | 1
| 1 | 2 | 14.302 | 16.7 | [18 empty] | 8
| 1 | 3 | 14.345 | 17 | [18 empty] | 16
| 1 | 4 | 14.348 | 17 | [18 empty] | 15.6751
| 1 | 5 | 14.389 | 17.2 | [18 empty] | 14
| 1 | 6 | 14.375 | 17.1
| Mean |  | 14.3502 | 17
| Median |  | 14.3465 | 17
| Std Dev. |  | 0.030142 | 0.16733
| Rel.Std.Dev. |  | 0.21% | 0.9843%
| PDM3 |  | -8.452% | 8.452%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for TiO2 in OREAS 555b (Certified Value 0.789 wt.%) | [21 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 23 | 25 | 27 | 28 |  |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.53 | 0.81 | 0.79 | 0.8 | 0.81 | 0.74 | 0.79 | 0.81 | 0.81 | 0.78 | 0.77589 | 0.7 | 0.79 | 0.81 | 1.41 | 0.688 | 0.84 | 0.8 |  |  | 1
| 1 | 2 | 0.7 | 0.82 | 0.79 | 0.79 | 0.81 | 0.75 | 0.78 | 0.82 | 0.83 | 0.8 | 0.77168 | 0.7 | 0.79 | 0.81 | 1.4 | 0.72 | 0.83 | 0.82 |  |  | 36
| 1 | 3 | 0.75 | 0.81 | 0.8 | 0.78 | 0.81 | 0.75 | 0.79 | 0.82 | 0.83 | 0.79 | 0.76045 | 0.7 | 0.79 | 0.79 | 1.39 | 0.7 | 0.83 | 0.83 |  |  | 16
| 1 | 4 | 0.67 | 0.81 | 0.79 | 0.79 | 0.81 | 0.74 | 0.78 | 0.82 | 0.82 | 0.79 | 0.77262 | 0.7 | 0.78 | 0.81 | 1.39 | 0.746 | 0.84 | 0.82 |  |  | 0.78852
| 1 | 5 | 0.64 | 0.82 | 0.8 | 0.79 | 0.81 | 0.76 | 0.79 | 0.83 | 0.81 | 0.78 | 0.76794 | 0.7 | 0.79 | 0.78 | 1.4 | 0.796 | 0.83 | 0.83 |  |  | 21
| 1 | 6 | 0.68 | 0.81 | 0.79 | 0.78 | 0.82 | 0.77 | 0.79 | 0.82 | 0.81 | 0.79 | 0.76607 | 0.7 | 0.79 | 0.79 | 1.41 | 0.753 | 0.84 | 0.82
| Mean |  | 0.66167 | 0.81333 | 0.79333 | 0.78833 | 0.81167 | 0.75167 | 0.78667 | 0.82 | 0.81833 | 0.78833 | 0.76911 | 0.7 | 0.78833 | 0.79833 | 1.4 | 0.73383 | 0.835 | 0.82
| Median |  | 0.675 | 0.81 | 0.79 | 0.79 | 0.81 | 0.75 | 0.79 | 0.82 | 0.815 | 0.79 | 0.76981 | 0.7 | 0.79 | 0.8 | 1.4 | 0.733 | 0.835 | 0.82
| Std Dev. |  | 0.07414 | 0.005164 | 0.005164 | 0.0075277 | 0.0040825 | 0.01169 | 0.005164 | 0.0063246 | 0.0098319 | 0.0075277 | 0.0054881 | 1.2162e-16 | 0.0040825 | 0.013292 | 0.0089443 | 0.039539 | 0.0054772 | 0.010954
| Rel.Std.Dev. |  | 11.2% | 0.6349% | 0.6509% | 0.9549% | 0.503% | 1.555% | 0.6564% | 0.7713% | 1.201% | 0.9549% | 0.7136% | 1.737e-14% | 0.5179% | 1.665% | 0.6389% | 5.388% | 0.656% | 1.336%
| PDM3 |  | -16.09% | 3.147% | 0.6108% | -0.02331% | 2.936% | -4.673% | -0.2347% | 3.993% | 3.781% | -0.02331% | -2.462% | -11.23% | -0.02331% | 1.245% | 77.55% | -6.935% | 5.895% | 3.993%
| Z-Score (Absolute) |  | 3.48 | 0.61 | 0.07 | 0.07 | 0.56 | 1.06 | 0.11 | 0.79 | 0.74 | 0.07 | 0.59 | 2.45 | 0.07 | 0.2 | 16.43 | 1.54 | 1.19 | 0.79
| Analytical results for Tl in OREAS 555b (Indicative Value 0.8 ppm) | [21 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 01 | 25 | [18 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | < 100 | 0.79 | [18 empty] | 1
| 1 | 2 | < 100 | 0.77 | [18 empty] | 26
| 1 | 3 | < 100 | 0.79 | [18 empty] | 16
| 1 | 4 | < 100 | 0.81 | [18 empty] | 0.79667
| 1 | 5 | < 100 | 0.82 | [18 empty] | 15
| 1 | 6 | < 100 | 0.8
| Mean |  |  | 0.79667
| Median |  |  | 0.795
| Std Dev. |  |  | 0.017512
| Rel.Std.Dev. |  |  | 2.198%
| PDM3 |  |  | -3.331e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Tm in OREAS 555b (Indicative Value 0.57 ppm) | [21 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.544 | 0.6 | [18 empty] | 1
| 1 | 2 | 0.533 | 0.6 | [18 empty] | 10
| 1 | 3 | 0.542 | 0.61 | [18 empty] | 16
| 1 | 4 | 0.543 | 0.6 | [18 empty] | 0.56608
| 1 | 5 | 0.536 | 0.55 | [18 empty] | 16
| 1 | 6 | 0.535 | 0.6
| Mean |  | 0.53883 | 0.59333
| Median |  | 0.539 | 0.6
| Std Dev. |  | 0.0047081 | 0.021602
| Rel.Std.Dev. |  | 0.8738% | 3.641%
| PDM3 |  | -4.814% | 4.814%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 555b (Indicative Value 9.13 ppm) | [21 empty] | Indicative
U | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 8.495 | 9.93 | [18 empty] | 1
| 1 | 2 | 8.24 | 9.57 | [18 empty] | 11
| 1 | 3 | 8.485 | 9.89 | [18 empty] | 16
| 1 | 4 | 8.486 | 9.81 | [18 empty] | 9.1254
| 1 | 5 | 8.482 | 9.83 | [18 empty] | 17
| 1 | 6 | 8.487 | 9.8
| Mean |  | 8.4458 | 9.805
| Median |  | 8.4855 | 9.82
| Std Dev. |  | 0.10093 | 0.1255
| Rel.Std.Dev. |  | 1.195% | 1.28%
| PDM3 |  | -7.447% | 7.447%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V2O5 in OREAS 555b (Certified Value 692 ppm) | [21 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 12 | 17 | 19 | 22 | 23 | 25 | 28 | [7 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 180 | 600 | 770 | 600 | 601.6 | 710 | 710 | 610 | < 1000 | 500 | 400 | 689.1 | 696 | [7 empty] | 1
| 1 | 2 | < 180 | 600 | 750 | 600 | 615.9 | 710 | 710 | 600 | < 1000 | 500 | 400 | 678.4 | 714 | [7 empty] | #N/A
| 1 | 3 | < 180 | 600 | 750 | 600 | 608.8 | 710 | 710 | 630 | < 1000 | 500 | 400 | 708.7 | 696 | [7 empty] | 16
| 1 | 4 | < 180 | 600 | 730 | 600 | 610.5 | 710 | 710 | 610 | < 1000 | 500 | 400 | 753.4 | 696 | [7 empty] | 691.7704
| 1 | 5 | < 180 | 600 | 770 | 700 | 633.7 | 710 | 710 | 610 | < 1000 | 500 | 400 | 817.6 | 732 | [7 empty] | 22
| 1 | 6 | 890 | 600 | 760 | 600 | 642.7 | 540 | 710 | 640 | < 1000 | 500 | 400 | 789.1 | 750
| Mean |  | 890 | 600 | 755 | 616.6667 | 618.8667 | 681.6667 | 710 | 616.6667 |  | 500 | 400 | 739.3833 | 714
| Median |  | 890 | 600 | 755 | 600 | 613.2 | 710 | 710 | 610 |  | 500 | 400 | 731.05 | 705
| Std Dev. |  |  | 0 | 15.1658 | 40.8248 | 15.9144 | 69.4022 | 0 | 15.0555 |  | 0 | 0 | 56.5242 | 22.7684
| Rel.Std.Dev. |  |  | 0% | 2.009% | 6.62% | 2.572% | 10.18% | 0% | 2.441% |  | 0% | 0% | 7.645% | 3.189%
| PDM3 |  | 28.66% | -13.27% | 9.14% | -10.86% | -10.54% | -1.461% | 2.635% | -10.86% |  | -27.72% | -42.18% | 6.883% | 3.213%
| Z-Score (Absolute) |  | NA | NA | 0.78 | NA | 1.13 | 0.21 | 0.21 | 1.16 | 2.8 | NA | NA | 0.56 | 0.21
Results from laboratories 1, 3, 6, 22 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 555b (Indicative Value 12.7 ppm) | [21 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 08 | 12 | 25 | [16 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [16 empty] | 1
1
| 1 | 1 | < 100 | 20 | < 10 | 3 | [16 empty] | 1
| 1 | 2 | < 100 | 10 | 20 | 3 | [16 empty] | 12
| 1 | 3 | < 100 | 20 | 30 | 2.9 | [16 empty] | 16
| 1 | 4 | < 100 | 20 | 40 | 3 | [16 empty] | 12.66
| 1 | 5 | < 100 | 20 | < 10 | 3.3 | [16 empty] | 18
| 1 | 6 | < 100 | 10 | 10 | 3
| Mean |  |  | 16.6667 | 25 | 3.0333
| Median |  |  | 20 | 25 | 3
| Std Dev. |  |  | 5.164 | 12.9099 | 0.13663
| Rel.Std.Dev. |  |  | 30.98% | 51.64% | 4.504%
| PDM3 |  |  | 31.65% | 97.47% | -76.04%
| Z-Score (Absolute) |  | 2.86 | 0.07 | 0.07 | 1.28
| Analytical results for Y in OREAS 555b (Indicative Value 34 ppm) | [21 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 12 | 13 | 21 | 25 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*MS | BF*OES | BF*XRF | [15 empty] | 1
1
| 1 | 1 | 39 | < 39 | 34.984 | 33 | 36.8 | [15 empty] | 1
| 1 | 2 | < 39 | 39 | 33.351 | 34 | 38.9 | [15 empty] | 13
| 1 | 3 | 39 | 39 | 34.526 | 34 | 38.8 | [15 empty] | 16
| 1 | 4 | < 39 | 63 | 34.368 | 35 | 38.8 | [15 empty] | 33.9574
| 1 | 5 | 47 | < 39 | 34.521 | 34 | 37.7 | [15 empty] | 19
| 1 | 6 | 39 | < 39 | 34.684 | 33 | 38.5
| Mean |  | 41 | 47 | 34.4057 | 33.8333 | 38.25
| Median |  | 39 | 39 | 34.5235 | 34 | 38.65
| Std Dev. |  | 4 | 13.8564 | 0.5574 | 0.75277 | 0.83606
| Rel.Std.Dev. |  | 9.756% | 29.48% | 1.62% | 2.225% | 2.186%
| PDM3 |  | 20.74% | 38.41% | 1.32% | -0.3654% | 12.64%
| Z-Score (Absolute) |  | 0 | 1.53 | 0.66 | 0.67 | 9.59
| Analytical results for Yb in OREAS 555b (Indicative Value 3.64 ppm) | [21 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 13 | 25 | [18 empty] | ppm
|  |  | BF*MS | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 3.631 | 3.7 | [18 empty] | 1
| 1 | 2 | 3.547 | 3.8 | [18 empty] | 14
| 1 | 3 | 3.584 | 3.6 | [18 empty] | 16
| 1 | 4 | 3.616 | 3.7 | [18 empty] | 3.6357
| 1 | 5 | 3.627 | 3.6 | [18 empty] | 20
| 1 | 6 | 3.623 | 3.6
| Mean |  | 3.6047 | 3.6667
| Median |  | 3.6195 | 3.65
| Std Dev. |  | 0.032892 | 0.08165
| Rel.Std.Dev. |  | 0.9125% | 2.227%
| PDM3 |  | -0.8527% | 0.8527%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 555b (Indicative Value 43.6 ppm) | [21 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 12 | 13 | 17 | 21 | 22 | 24 | 25 | 27 | 28 | [6 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 1
1
| 1 | 1 | < 100 | 1500 | < 80 | 44 | < 50 | < 50 | < 100 | < 50 | 30 | 640 | < 40 | 19 | 20 | 100 | [6 empty] | 1
| 1 | 2 | < 100 | 1500 | < 80 | 43 | < 50 | < 50 | < 100 | < 50 | 32 | 640 | < 40 | 18 | 20 | 200 | [6 empty] | 15
| 1 | 3 | < 100 | 1500 | < 80 | 42 | < 50 | < 50 | < 100 | < 50 | 33 | 640 | < 40 | 17 | 20 | 100 | [6 empty] | 16
| 1 | 4 | < 100 | 1500 | < 80 | 43 | 50 | < 50 | < 100 | < 50 | 29 | 640 | < 40 | 18 | 20 | 100 | [6 empty] | 43.6111
| 1 | 5 | < 100 | 1500 | < 80 | 45 | < 50 | < 50 | < 100 | < 50 | 32 | 560 | < 40 | 16 | 20 | 200 | [6 empty] | 21
| 1 | 6 | < 100 | 1500 | < 80 | 44 | < 50 | < 50 | < 100 | < 50 | 28 | 640 | < 40 | 17 | 20 | 100
| Mean |  |  | 1500 |  | 43.5 | 50 |  |  |  | 30.6667 | 626.6667 |  | 17.5 | 20 | 133.3333
| Median |  |  | 1500 |  | 43.5 | 50 |  |  |  | 31 | 640 |  | 17.5 | 20 | 100
| Std Dev. |  |  | 0 |  | 1.0488 |  |  |  |  | 1.9664 | 32.6599 |  | 1.0488 | 0 | 51.6398
| Rel.Std.Dev. |  |  | 0% |  | 2.411% |  |  |  |  | 6.412% | 5.212% |  | 5.993% | 0% | 38.73%
| PDM3 |  |  | 3339% |  | -0.2548% | 14.65% |  |  |  | -29.68% | 1337% |  | -59.87% | -54.14% | 205.7%
| Z-Score (Absolute) |  | 0.67 | 67.72 | 0.22 | 0.37 | 0.29 | 0.48 | 0.67 | 0.48 | 0.22 | 27.46 | 0.71 | 0.83 | 0.71 | 4.53
| Analytical results for Zr in OREAS 555b (Certified Value 238 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 06 | 07 | 08 | 12 | 21 | 22 | 25 | 28 | [11 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | 200 | 220 | 238 | 300 | 300 | 230 | < 70 | 238 | 300 | [11 empty] | 1
| 1 | 2 | 200 | 220 | 237 | 300 | 300 | 234 | 70 | 237 | 300 | [11 empty] | 18
| 1 | 3 | 300 | 220 | 234 | 300 | 200 | 233 | 70 | 236 | 300 | [11 empty] | 16
| 1 | 4 | 200 | 300 | 236 | 300 | 200 | 229 | < 70 | 244 | 300 | [11 empty] | 237.6548
| 1 | 5 | 300 | 300 | 238 | 300 | 300 | 239 | < 70 | 236 | 300 | [11 empty] | 23
| 1 | 6 | 200 | 220 | 238 | 300 | 300 | 235 | 70 | 237 | 200
| Mean |  | 233.3333 | 246.6667 | 236.8333 | 300 | 266.6667 | 233.3333 | 70 | 238 | 283.3333
| Median |  | 200 | 220 | 237.5 | 300 | 300 | 233.5 | 70 | 237 | 300
| Std Dev. |  | 51.6398 | 41.3118 | 1.6021 | 0 | 51.6398 | 3.6148 | 0 | 3.0332 | 40.8248
| Rel.Std.Dev. |  | 22.13% | 16.75% | 0.6765% | 0% | 19.36% | 1.549% | 0% | 1.274% | 14.41%
| PDM3 |  | -1.818% | 3.792% | -0.3457% | 26.23% | 12.21% | -1.818% | -70.55% | 0.1453% | 19.22%
| Z-Score (Absolute) |  | 0.51 | 3.38 | 0.51 | NA | NA | 0.51 | 51.89 | 0.84 | NA
Results from laboratories 8, 12 and 28 were removed due to their 100 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C813: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C831: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C867: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C885: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C903: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C921: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C939: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C957: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C975: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C993: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1011: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1029: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1048: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1066: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1084: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1102: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1120: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 555b (Certified Value 12.92 wt.%) | [25 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | wt.%
|  |  | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | BF*XRF | BF*XRF | BF*XRF | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 12.97 | 13.05 | 12.84 | 12.67 | 12.87 | 12.8 | 13.04 | 12.45 | 13.2 | 12.64 | 12.68 | 13.01 | 12.9 | 13.2154 | 13.7233 | 12.6684 | 12.6269 | 13.64 | 13.2 | 12.34 | 13.4159 | 15.5 | 12.6 | 12.48 | 1
| 1 | 2 | 12.89 | 13.09 | 12.88 | 12.63 | 12.85 | 12.8 | 13 | 12.41 | 13.24 | 12.76 | 12.61 | 12.95 | 13 | 13.1955 | 13.7069 | 12.725 | 12.6386 | 13.28 | 13.2 | 12.51 | 13.5706 | 13.26 | 12.6 | 12.51 | #N/A
| 1 | 3 | 12.98 | 13.04 | 12.88 | 12.61 | 12.86 | 12.9 | 13.14 | 12.4 | 13.35 | 12.72 | 12.55 | 12.99 | 12.9 | 13.1133 | 13.8151 | 12.5525 | 12.6277 | 13.35 | 13.3 | 12.53 | 13.6734 | 13.12 | 12.6 | 12.62 | 16
| 1 | 4 | 12.97 | 13.01 | 12.87 | 12.66 | 12.86 | 12.8 | 13.06 | 12.38 | 13.23 | 12.67 | 12.62 | 12.98 | 12.8 | 13.096 | 13.8292 | 12.6014 | 12.9483 | 13.3 | 13.2 | 12.16 | 13.4149 | 13.07 | 12.6 | 12.55 | 12.9193
| 1 | 5 | 12.79 | 13.15 | 12.89 | 12.75 | 12.83 | 12.8 | 13.05 | 12.51 | 13.18 | 12.66 | 12.61 | 12.97 | 12.8 | 13.106 | 13.7855 | 12.6143 | 12.6132 | 13.24 | 13.2 | 12.53 | 13.6093 | 13.07 | 12.6 | 12.51 | 25
| 1 | 6 | 12.85 | 13.16 | 12.91 | 12.82 | 12.86 | 12.8 | 13.05 | 12.38 | 13.28 | 12.66 | 10.83 | 12.97 | 12.9 | 13.1397 | 13.7359 | 12.5959 | 12.6515 | 13.19 | 13.3 | 12.4 | 13.3367 | 13.01 | 12.6 | 12.55
| Mean |  | 12.9083 | 13.0833 | 12.8783 | 12.69 | 12.855 | 12.8167 | 13.0567 | 12.4217 | 13.2467 | 12.685 | 12.3167 | 12.9783 | 12.8833 | 13.1443 | 13.766 | 12.6262 | 12.6844 | 13.3333 | 13.2333 | 12.4117 | 13.5035 | 13.505 | 12.6 | 12.5367
| Median |  | 12.93 | 13.07 | 12.88 | 12.665 | 12.86 | 12.8 | 13.05 | 12.405 | 13.235 | 12.665 | 12.61 | 12.975 | 12.9 | 13.1265 | 13.7607 | 12.6079 | 12.6331 | 13.29 | 13.2 | 12.455 | 13.4932 | 13.095 | 12.6 | 12.53
| Std Dev. |  | 0.078081 | 0.06121 | 0.023166 | 0.07975 | 0.013784 | 0.040825 | 0.045898 | 0.050365 | 0.06121 | 0.045497 | 0.72948 | 0.020412 | 0.075277 | 0.049927 | 0.051004 | 0.061075 | 0.12994 | 0.15971 | 0.05164 | 0.1458 | 0.13263 | 0.98099 | 0 | 0.048854
| Rel.Std.Dev. |  | 0.6049% | 0.4678% | 0.1799% | 0.6284% | 0.1072% | 0.3185% | 0.3515% | 0.4055% | 0.4621% | 0.3587% | 5.923% | 0.1573% | 0.5843% | 0.3798% | 0.3705% | 0.4837% | 1.024% | 1.198% | 0.3902% | 1.175% | 0.9822% | 7.264% | 0% | 0.3897%
| PDM3 |  | -0.08473% | 1.27% | -0.3169% | -1.775% | -0.4976% | -0.7943% | 1.063% | -3.852% | 2.534% | -1.813% | -4.664% | 0.4571% | -0.2782% | 1.742% | 6.554% | -2.268% | -1.818% | 3.205% | 2.431% | -3.929% | 4.522% | 4.534% | -2.471% | -2.962%
| Z-Score (Absolute) |  | 0.07 | 0.53 | 0.01 | 0.5 | 0.07 | 0.17 | 0.46 | 1.2 | 0.95 | 0.51 | 1.47 | 0.25 | 0.01 | 0.69 | 2.3 | 0.66 | 0.51 | 1.18 | 0.92 | 1.22 | 1.62 | 1.62 | 0.73 | 0.9
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 555b (Certified Value 4.53 wt.%) | [27 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC |  | 2
3
| 1 | 1 | 4.33 | 4.53 | 4.33 | 4.44 | 4.94 | 4.68 | 4.2921 | 4.43 | 4.37 | 4.53 | 4.47 | 4.49 | 4.85 | 3.9285 | 4.81 | 4.45 | 4.51 | 4.42 | 4.44 | 4.65 | 4.68 | 4.86 | 4.7 | 4.19 | 4.48 |  | 1
| 1 | 2 | 4.24 | 4.63 | 4.29 | 4.46 | 4.78 | 4.55 | 4.2836 | 4.48 | 4.29 | 4.51 | 4.5 | 4.48 | 4.9 | 3.9471 | 4.79 | 4.47 | 4.48 | 4.42 | 4.44 | 4.63 | 4.76 | 4.91 | 4.8 | 4.2 | 4.48 |  | #N/A
| 1 | 3 | 4.22 | 4.58 | 4.33 | 4.48 | 4.7 | 4.67 | 4.2424 | 4.44 | 4.39 | 4.52 | 4.42 | 4.5 | 4.89 | 3.9176 | 4.77 | 4.48 | 4.47 | 4.47 | 4.46 | 4.64 | 4.67 | 4.82 | 4.65 | 4.2 | 4.44 |  | 16
| 1 | 4 | 4.24 | 4.6 | 4.35 | 4.5 | 4.75 | 4.61 | 4.2806 | 4.45 | 4.41 | 4.59 | 4.44 | 4.48 | 4.93 | 3.9122 | 4.81 | 4.46 | 4.51 | 4.48 | 4.45 | 4.65 | 4.83 | 5.01 | 4.69 | 4.17 | 4.51 |  | 4.5302
| 1 | 5 | 4.27 | 4.56 | 4.28 | 4.5 | 4.7 | 4.5 | 4.3369 | 4.43 | 4.44 | 4.56 | 4.44 | 4.5 | 4.9 | 3.9481 | 4.7 | 4.46 | 4.51 | 4.5 | 4.47 | 4.65 | 4.72 | 4.71 | 4.65 | 4.2 | 4.37 |  | 27
| 1 | 6 | 4.3 | 4.53 | 4.34 | 4.49 | 4.73 | 4.58 | 4.3576 | 4.43 | 4.44 | 4.48 | 4.38 | 4.47 | 4.92 | 3.9365 | 4.8 | 4.48 | 4.52 | 4.48 | 4.5 | 4.65 | 4.77 | 4.92 | 4.65 | 4.2 | 4.49
| Mean |  | 4.2667 | 4.5717 | 4.32 | 4.4783 | 4.7667 | 4.5983 | 4.2989 | 4.4433 | 4.39 | 4.5317 | 4.4417 | 4.4867 | 4.8983 | 3.9317 | 4.78 | 4.4667 | 4.5 | 4.4617 | 4.46 | 4.645 | 4.7383 | 4.8717 | 4.69 | 4.1933 | 4.4617
| Median |  | 4.255 | 4.57 | 4.33 | 4.485 | 4.74 | 4.595 | 4.2878 | 4.435 | 4.4 | 4.525 | 4.44 | 4.485 | 4.9 | 3.9325 | 4.795 | 4.465 | 4.51 | 4.475 | 4.455 | 4.65 | 4.74 | 4.885 | 4.67 | 4.2 | 4.48
| Std Dev. |  | 0.041793 | 0.039707 | 0.028284 | 0.024014 | 0.090259 | 0.06969 | 0.041706 | 0.019664 | 0.056214 | 0.038687 | 0.041191 | 0.012111 | 0.027869 | 0.014949 | 0.041952 | 0.012111 | 0.02 | 0.033714 | 0.022804 | 0.0083666 | 0.06047 | 0.10187 | 0.05831 | 0.012111 | 0.050365
| Rel.Std.Dev. |  | 0.9795% | 0.8686% | 0.6547% | 0.5362% | 1.894% | 1.516% | 0.9702% | 0.4425% | 1.28% | 0.8537% | 0.9274% | 0.2699% | 0.5689% | 0.3802% | 0.8777% | 0.2711% | 0.4444% | 0.7556% | 0.5113% | 0.1801% | 1.276% | 2.091% | 1.243% | 0.2888% | 1.129%
| PDM3 |  | -5.818% | 0.9143% | -4.641% | -1.146% | 5.219% | 1.503% | -5.107% | -1.918% | -3.096% | 0.03138% | -1.955% | -0.9619% | 8.125% | -13.21% | 5.513% | -1.403% | -0.6676% | -1.514% | -1.551% | 2.533% | 4.593% | 7.536% | 3.526% | -7.437% | -1.514%
| Z-Score (Absolute) |  | 1.19 | 0.52 | 0.89 | 0 | 1.62 | 0.67 | 1.01 | 0.2 | 0.5 | 0.3 | 0.21 | 0.05 | 2.36 | 3.07 | 1.7 | 0.07 | 0.12 | 0.09 | 0.1 | 0.94 | 1.46 | 2.21 | 1.19 | 1.6 | 0.09
| Analytical results for S in OREAS 555b (Certified Value 4.48 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 4.69 | 4.51 | 4.36 | 4.49 | 4.67 | 4.74 | 4.5857 | 4.45 | 4.44 | 4.4 | 4.44 | 4.53 | 4.43 | 4.5 | 4.5235 | 4.48 | 4.51 | 4.41 | 4.45 | 4.45 | 4.45 | 5.05 | 4.69 | 4.34 | 4.69 | 4.41 | 1
| 1 | 2 | 4.65 | 4.48 | 4.38 | 4.48 | 4.56 | 4.72 | 4.6008 | 4.36 | 4.64 | 4.45 | 4.39 | 4.44 | 4.42 | 4.5 | 4.5216 | 4.47 | 4.52 | 4.41 | 4.43 | 4.47 | 4.41 | 5.07 | 4.54 | 4.39 | 4.69 | 4.42 | 28
| 1 | 3 | 4.67 | 4.43 | 4.36 | 4.49 | 4.52 | 4.74 | 4.6221 | 4.54 | 4.63 | 4.43 | 4.39 | 4.47 | 4.45 | 4.5 | 4.5446 | 4.49 | 4.53 | 4.28 | 4.43 | 4.41 | 4.42 | 5.06 | 4.61 | 4.34 | 4.56 | 4.38 | 16
| 1 | 4 | 4.65 | 4.51 | 4.38 | 4.53 | 4.48 | 4.72 | 4.4993 | 4.52 | 4.69 | 4.42 | 4.39 | 4.46 | 4.41 | 4.5 | 4.5401 | 4.49 | 4.52 | 4.28 | 4.45 | 4.38 | 4.44 | 5.05 | 4.54 | 4.32 | 4.62 | 4.45 | 4.4824
| 1 | 5 | 4.68 | 4.41 | 4.35 | 4.55 | 4.49 | 4.72 | 4.4876 | 4.45 | 4.61 | 4.48 | 4.44 | 4.49 | 4.43 | 4.5 | 4.5592 | 4.47 | 4.54 | 4.43 | 4.44 | 4.42 | 4.43 | 5.06 | 4.38 | 4.3 | 4.67 | 4.39 | 28
| 1 | 6 | 4.64 | 4.28 | 4.39 | 4.53 | 4.5 | 4.72 | 4.6386 | 4.48 | 4.57 | 4.46 | 4.27 | 4.5 | 4.44 | 4.6 | 4.5599 | 4.49 | 4.5 | 4.41 | 4.45 | 4.42 | 4.43 | 5.09 | 4.63 | 4.3 | 4.37 | 4.42
| Mean |  | 4.6633 | 4.4367 | 4.37 | 4.5117 | 4.5367 | 4.7267 | 4.5723 | 4.4667 | 4.5967 | 4.44 | 4.3867 | 4.4817 | 4.43 | 4.5167 | 4.5415 | 4.4817 | 4.52 | 4.37 | 4.4417 | 4.425 | 4.43 | 5.0633 | 4.565 | 4.3317 | 4.6 | 4.4117
| Median |  | 4.66 | 4.455 | 4.37 | 4.51 | 4.51 | 4.72 | 4.5932 | 4.465 | 4.62 | 4.44 | 4.39 | 4.48 | 4.43 | 4.5 | 4.5424 | 4.485 | 4.52 | 4.41 | 4.445 | 4.42 | 4.43 | 5.06 | 4.575 | 4.33 | 4.645 | 4.415
| Std Dev. |  | 0.019664 | 0.087101 | 0.015492 | 0.028577 | 0.071181 | 0.010328 | 0.063829 | 0.06377 | 0.086178 | 0.028983 | 0.062183 | 0.031885 | 0.014142 | 0.040825 | 0.016629 | 0.0098319 | 0.014142 | 0.070143 | 0.0098319 | 0.031464 | 0.014142 | 0.015055 | 0.1071 | 0.033714 | 0.12329 | 0.024833
| Rel.Std.Dev. |  | 0.4217% | 1.963% | 0.3545% | 0.6334% | 1.569% | 0.2185% | 1.396% | 1.428% | 1.875% | 0.6528% | 1.418% | 0.7115% | 0.3192% | 0.9039% | 0.3661% | 0.2194% | 0.3129% | 1.605% | 0.2214% | 0.7111% | 0.3192% | 0.2973% | 2.346% | 0.7783% | 2.68% | 0.5629%
| PDM3 |  | 4.036% | -1.021% | -2.508% | 0.6527% | 1.21% | 5.449% | 2.006% | -0.3512% | 2.549% | -0.9462% | -2.136% | -0.0166% | -1.169% | 0.7642% | 1.318% | -0.0166% | 0.8386% | -2.508% | -0.909% | -1.281% | -1.169% | 12.96% | 1.843% | -3.363% | 2.623% | -1.578%
| Z-Score (Absolute) |  | 2.1 | 0.52 | 1.29 | 0.35 | 0.64 | 2.84 | 1.05 | 0.17 | 1.33 | 0.48 | 1.1 | 0 | 0.6 | 0.41 | 0.69 | 0 | 0.44 | 1.29 | 0.46 | 0.66 | 0.6 | 6.73 | 0.96 | 1.74 | 1.37 | 0.81
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach (05%)
| Analytical results for Co in OREAS 555b (Certified Value 0.432 wt.%) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 13 | 15 | 18 | 19 | 23 | 25 | [10 empty] | wt.%
|  |  | AL-5S*AAS | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | [10 empty] | 3
3
| 1 | 1 | 0.4237 | 0.4562 | 0.456 | 0.36 | 0.43235 | 0.45408 | 0.6625 | 0.414 | 0.3655 | [10 empty] | 1
| 1 | 2 | 0.4397 | 0.4562 | 0.463 | 0.38 | 0.43862 | 0.49119 | 0.7367 | 0.42444 | 0.3774 | [10 empty] | 36
| 1 | 3 | 0.4311 | 0.4779 | 0.481 | 0.36 | 0.42905 | 0.48837 | 0.6899 | 0.42502 | 0.4031 | [10 empty] | 16
| 1 | 4 | 0.4331 | 0.4621 | 0.443 | 0.37 | 0.44834 | 0.45261 | 0.7826 | 0.41823 | 0.3878 | [10 empty] | 0.43156
| 1 | 5 | 0.4345 | 0.4615 | 0.468 | 0.41 | 0.42299 | 0.47499 | 0.7484 | 0.42616 | 0.3803 | [10 empty] | 30
| 1 | 6 | 0.4376 | 0.4395 | 0.49 | 0.4 | 0.44646 | 0.47864 | 0.4959 | 0.41565 | 0.3857
| Mean |  | 0.43328 | 0.4589 | 0.46683 | 0.38 | 0.4363 | 0.47331 | 0.686 | 0.42058 | 0.3833
| Median |  | 0.4338 | 0.45885 | 0.4655 | 0.375 | 0.43548 | 0.47681 | 0.7133 | 0.42133 | 0.383
| Std Dev. |  | 0.0056201 | 0.012401 | 0.016964 | 0.020976 | 0.0099905 | 0.016588 | 0.10247 | 0.0052693 | 0.01248
| Rel.Std.Dev. |  | 1.297% | 2.702% | 3.634% | 5.52% | 2.29% | 3.505% | 14.94% | 1.253% | 3.256%
| PDM3 |  | 0.3983% | 6.334% | 8.172% | -11.95% | 1.098% | 9.674% | 58.96% | -2.544% | -11.18%
| Z-Score (Absolute) |  | 0.07 | 0.5 | 0.67 | 1.24 | 0 | 0.82 | 5.51 | 0.35 | 1.17
| Analytical results for Cu in OREAS 555b (Certified Value 1.19 wt.%) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 08 | 09 | 10 | 11 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 25 | 27 | 28 | wt.%
|  |  | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | 2
3
| 1 | 1 | 1.175 | 1.4665 | 1.2632 | 1.2 | 1.12 | 1.17 | 1.06 | 1.15 | 1.32 | 1.1619 | 1.28 | 1.3639 | 1.0622 | 1.214 | 1.226 | 0.99474 | 0.9253 | 1.189 | 1.192 | 1
| 1 | 2 | 1.187 | 1.4835 | 1.2446 | 1.21 | 1.11 | 1.155 | 1.18 | 1.125 | 1.24 | 1.1921 | 1.2 | 1.3575 | 1.0946 | 1.222 | 1.21 | 0.99574 | 0.9487 | 1.19 | 1.2 | 39
| 1 | 3 | 1.174 | 1.4878 | 1.2815 | 1.2 | 1.125 | 1.165 | 1.165 | 1.19 | 1.29 | 1.1852 | 1.19 | 1.3611 | 1.1054 | 1.219 | 1.222 | 0.96981 | 1.004 | 1.198 | 1.118 | 16
| 1 | 4 | 1.168 | 1.479 | 1.265 | 1.18 | 1.11 | 1.165 | 1.145 | 1.14 | 1.31 | 1.1408 | 1.22 | 1.3529 | 1.1506 | 1.218 | 1.212 | 0.97996 | 0.955 | 1.21 | 1.17 | 1.1872
| 1 | 5 | 1.176 | 1.4774 | 1.2472 | 1.21 | 1.12 | 1.165 | 1.165 | 1.165 | 1.32 | 1.147 | 1.2 | 1.3497 | 1.1174 | 1.229 | 1.225 | 0.98123 | 0.9833 | 1.198 | 1.153 | 31
| 1 | 6 | 1.161 | 1.478 | 1.2158 | 1.22 | 1.105 | 1.14 | 1.17 | 1.135 | 1.33 | 1.2052 | 1.19 | 1.3592 | 0.9293 | 1.224 | 1.216 | 0.97433 | 0.9532 | 1.168 | 1.149
| Mean |  | 1.1735 | 1.4787 | 1.2529 | 1.2033 | 1.115 | 1.16 | 1.1475 | 1.1508 | 1.3017 | 1.172 | 1.2133 | 1.3574 | 1.0766 | 1.221 | 1.2185 | 0.98263 | 0.96158 | 1.1922 | 1.1637
| Median |  | 1.1745 | 1.4785 | 1.2552 | 1.205 | 1.115 | 1.165 | 1.165 | 1.145 | 1.315 | 1.1736 | 1.2 | 1.3583 | 1.1 | 1.2205 | 1.219 | 0.9806 | 0.9541 | 1.194 | 1.1615
| Std Dev. |  | 0.0086891 | 0.0071688 | 0.022575 | 0.013663 | 0.007746 | 0.010954 | 0.044357 | 0.023541 | 0.033116 | 0.026018 | 0.034448 | 0.0052723 | 0.077716 | 0.0052154 | 0.0068044 | 0.010592 | 0.02781 | 0.014034 | 0.030257
| Rel.Std.Dev. |  | 0.7404% | 0.4848% | 1.802% | 1.135% | 0.6947% | 0.9443% | 3.865% | 2.046% | 2.544% | 2.22% | 2.839% | 0.3884% | 7.219% | 0.4271% | 0.5584% | 1.078% | 2.892% | 1.177% | 2.6%
| PDM3 |  | -1.154% | 24.55% | 5.533% | 1.359% | -6.081% | -2.291% | -3.344% | -3.063% | 9.642% | -1.278% | 2.202% | 14.34% | -9.317% | 2.847% | 2.637% | -17.23% | -19% | 0.4186% | -1.982%
| Z-Score (Absolute) |  | 0 | 4.57 | 1.19 | 0.45 | 0.88 | 0.2 | 0.39 | 0.34 | 1.92 | 0.02 | 0.6 | 2.76 | 1.45 | 0.71 | 0.67 | 2.86 | 3.18 | 0.28 | 0.15
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 555b (Indicative Value 0.426 ppm) | [23 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 25 | 26 | [19 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | [19 empty] | 3
3
| 1 | 1 | 0.79 | 0.1 | < 5 | [19 empty] | 1
| 1 | 2 | 0.8 | 0.1 | < 5 | [19 empty] | 1
| 1 | 3 | 0.77 | 0.1 | < 5 | [19 empty] | 16
| 1 | 4 | 0.72 | < 0.1 | < 5 | [19 empty] | 0.42583
| 1 | 5 | 0.73 | 0.1 | < 5 | [19 empty] | 23
| 1 | 6 | 0.75 | 0.1 | < 5
| Mean |  | 0.76 | 0.1
| Median |  | 0.76 | 0.1
| Std Dev. |  | 0.032249 | 0
| Rel.Std.Dev. |  | 4.243% | 0%
| PDM3 |  | 78.47% | -76.52%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.76
| Analytical results for Al in OREAS 555b (Certified Value 6.42 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 6.09 | 6.46 | 6.37 | 6.76 | 6.43 | 6.51 | 6.377 | 6.219 | 6.245 | 6.563 | 6.589 | 6.7539 | 8.1113 | 6.67 | 6.8071 | 6.1223 | 6.4 | 6.81 | 6.02 | 6.36 | 6.32 | 6.58 | 1
| 1 | 2 | 6 | 6.35 | 6.38 | 6.76 | 6.72 | 6.52 | 6.43 | 6.298 | 5.98 | 6.457 | 6.483 | 6.6274 | 8.1566 | 6.78 | 6.4921 | 6.3885 | 6.45 | 7.19 | 5.04 | 6.53 | 6.31 | 6.7 | #N/A
| 1 | 3 | 5.83 | 6.45 | 6.39 | 6.71 | 6.49 | 6.57 | 6.404 | 6.245 | 6.113 | 6.43 | 6.536 | 6.436 | 8.1299 | 6.61 | 6.4412 | 6.2259 | 6.43 | 6.91 | 4.3 | 6.33 | 6.23 | 6.54 | 16
| 1 | 4 | 6.06 | 6.5 | 6.33 | 6.73 | 6.43 | 6.54 | 6.404 | 6.324 | 6.43 | 6.272 | 6.51 | 6.6817 | 8.1854 | 6.6 | 6.4802 | 6.4019 | 6.37 | 6.88 | 2.79 | 6.45 | 6.25 | 6.6 | 6.4231
| 1 | 5 | 5.99 | 6.62 | 6.37 | 6.43 | 6.38 | 6.53 | 6.404 | 6.219 | 5.901 | 6.351 | 6.51 | 6.4824 | 8.1328 | 6.67 | 6.4709 | 6.3605 | 6.41 | 7.13 | 4.77 | 6.38 | 6.26 | 6.48 | 33
| 1 | 6 | 5.95 | 6.54 | 6.38 | 6.66 | 6.49 | 6.51 | 6.404 | 6.272 | 6.245 | 6.457 | 6.536 | 6.4603 | 8.1305 | 6.74 | 6.4673 | 6.337 | 6.44 | 7.19 | 3.81 | 6.33 | 6.28 | 6.59
| Mean |  | 5.9867 | 6.4867 | 6.37 | 6.675 | 6.49 | 6.53 | 6.4038 | 6.2628 | 6.1523 | 6.4217 | 6.5273 | 6.5736 | 8.1411 | 6.6783 | 6.5265 | 6.306 | 6.4167 | 7.0183 | 4.455 | 6.3967 | 6.275 | 6.5817
| Median |  | 5.995 | 6.48 | 6.375 | 6.72 | 6.46 | 6.525 | 6.404 | 6.2585 | 6.179 | 6.4435 | 6.523 | 6.5549 | 8.1317 | 6.67 | 6.4755 | 6.3487 | 6.42 | 7.02 | 4.535 | 6.37 | 6.27 | 6.585
| Std Dev. |  | 0.091797 | 0.091141 | 0.020976 | 0.12566 | 0.12017 | 0.022804 | 0.016762 | 0.042948 | 0.19426 | 0.099935 | 0.036109 | 0.13206 | 0.026067 | 0.070828 | 0.13853 | 0.10963 | 0.029439 | 0.17069 | 1.1048 | 0.078909 | 0.035071 | 0.072778
| Rel.Std.Dev. |  | 1.533% | 1.405% | 0.3293% | 1.883% | 1.852% | 0.3492% | 0.2618% | 0.6858% | 3.158% | 1.556% | 0.5532% | 2.009% | 0.3202% | 1.061% | 2.123% | 1.739% | 0.4588% | 2.432% | 24.8% | 1.234% | 0.5589% | 1.106%
| PDM3 |  | -6.795% | 0.9899% | -0.8265% | 3.922% | 1.042% | 1.665% | -0.2997% | -2.495% | -4.215% | -0.0221% | 1.623% | 2.344% | 26.75% | 3.974% | 1.61% | -1.823% | -0.09994% | 9.267% | -30.64% | -0.4113% | -2.306% | 2.469%
| Z-Score (Absolute) |  | 2.55 | 0.18 | 0.46 | 1.21 | 0.2 | 0.41 | 0.27 | 1.04 | 1.65 | 0.18 | 0.4 | 0.65 | 9.21 | 1.22 | 0.4 | 0.81 | 0.2 | 3.08 | 10.91 | 0.31 | 0.98 | 0.7
| Analytical results for As in OREAS 555b (Certified Value 66 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 21 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 0
1
| 1 | 1 | 45.5 | < 100 | 76.7 | 68 | 70 | 65 | 100 | 100 | 100 | 200 | 100 | 63.2611 | 70 | 60 | 45 | 63 | 64 | < 100 | < 100 |  |  |  | 1
| 1 | 2 | 41.6 | < 100 | 75.8 | 70 | 62 | 68 | 100 | 100 | 100 | 200 | 100 | 63.687 | 70 | 70 | 48 | 56 | 56 | < 100 | < 100 |  |  |  | 2
| 1 | 3 | 45.3 | < 100 | 76.4 | 67 | 65 | 67 | 100 | 100 | 100 | 200 | 100 | 63.2961 | 80 | 70 | 48 | 66 | 65 | < 100 | < 100 |  |  |  | 16
| 1 | 4 | 42 | < 100 | 76.3 | 68 | 59 | 66 | 100 | 100 | 100 | 200 | 100 | 63.5098 | 80 | 60 | 44 | 56 | 57 | < 100 | < 100 |  |  |  | 65.7417
| 1 | 5 | 41.8 | < 100 | 76.1 | 70 | 60 | 64 | 100 | 100 | 100 | 300 | 200 | 63.6683 | 70 | 70 | 44 | 60 | 62 | < 100 | < 100 |  |  |  | 34
| 1 | 6 | 44.3 | < 100 | 76.9 | 69 | 70 | 64 | 100 | 100 | 100 | 200 | 100 | 63.5282 | 80 | 70 | 47 | 59 | 66 | < 100 | < 100
| Mean |  | 43.4167 |  | 76.3667 | 68.6667 | 64.3333 | 65.6667 | 100 | 100 | 100 | 216.6667 | 116.6667 | 63.4917 | 75 | 66.6667 | 46 | 60 | 61.6667
| Median |  | 43.15 |  | 76.35 | 68.5 | 63.5 | 65.5 | 100 | 100 | 100 | 200 | 100 | 63.519 | 75 | 70 | 46 | 59.5 | 63
| Std Dev. |  | 1.8214 |  | 0.39833 | 1.2111 | 4.8442 | 1.633 | 0 | 0 | 0 | 40.8248 | 40.8248 | 0.18022 | 5.4772 | 5.164 | 1.8974 | 3.9497 | 4.2269
| Rel.Std.Dev. |  | 4.195% |  | 0.5216% | 1.764% | 7.53% | 2.487% | 0% | 0% | 0% | 18.84% | 34.99% | 0.2838% | 7.303% | 7.746% | 4.125% | 6.583% | 6.854%
| PDM3 |  | -33.96% |  | 16.16% | 4.449% | -2.142% | -0.1141% | 52.11% | 52.11% | 52.11% | 229.6% | 77.46% | -3.422% | 14.08% | 1.407% | -30.03% | -8.734% | -6.199%
| Z-Score (Absolute) |  | 1.26 | 0.78 | 1.12 | 0.57 | 0.25 | 0.35 | NA | NA | NA | NA | NA | 0.19 | NA | NA | 1.07 | 0.06 | 0.06 | 0.78 | 0.78
Results from laboratories 17 and 21 were removed due to their 10 ppm reading resolution. Results from laboratories 8, 9, 10, 11 and 12 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 555b (Certified Value 232 ppm) | [23 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 19 | 21 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | [14 empty] | 0
0
| 1 | 1 | 240 | 270 | 228 | 236.8074 | 236 | 210 | 123 | 199 | [14 empty] | 1
| 1 | 2 | 239 | 290 | 232 | 255.8596 | 246 | 200 | 145 | 205 | [14 empty] | #N/A
| 1 | 3 | 242 | 280 | 230 | 235.3208 | 233 | 200 | 223 | 203 | [14 empty] | 16
| 1 | 4 | 240 | 260 | 232 | 262.9591 | 241 | 200 | 273 | 200 | [14 empty] | 231.8282
| 1 | 5 | 229 | 270 | 221 | 263.1601 | 237 | 200 | 257 | 200 | [14 empty] | 35
| 1 | 6 | 233 | 290 | 224 | 261.8446 | 237 | 210 | 191 | 198
| Mean |  | 237.1667 | 276.6667 | 227.8333 | 252.6586 | 238.3333 | 203.3333 | 202 | 200.8333
| Median |  | 239.5 | 275 | 229 | 258.8521 | 237 | 200 | 207 | 200
| Std Dev. |  | 5.0365 | 12.1106 | 4.4907 | 13.1367 | 4.5461 | 5.164 | 60.1963 | 2.6394
| Rel.Std.Dev. |  | 2.124% | 4.377% | 1.971% | 5.199% | 1.907% | 2.54% | 29.8% | 1.314%
| PDM3 |  | 2.303% | 19.34% | -1.723% | 8.985% | 2.806% | -12.29% | -12.87% | -13.37%
| Z-Score (Absolute) |  | 0.13 | 1.21 | 0.13 | 0.55 | 0.16 | 0.8 | 0.83 | 0.87
| Analytical results for Ba in OREAS 555b (Certified Value 259 ppm) | [23 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 266.8 | 280 | 239 | 259 | 264 | 275 | 262.9673 | 286 | 272.9884 | 258 | 240 | 254 | 212 | 256 | 264 | [7 empty] | 1
| 1 | 2 | 265.8 | 271 | 236 | 260 | 260 | 272 | 261.4677 | 287 | 262.6257 | 275 | 250 | 264 | 208 | 255 | 251 | [7 empty] | 30
| 1 | 3 | 266.8 | 267 | 238 | 258 | 260 | 276 | 248.4895 | 285 | 247.3475 | 262 | 240 | 244 | 178 | 247 | 260 | [7 empty] | 16
| 1 | 4 | 270.4 | 270 | 233 | 266 | 266 | 281 | 262.1204 | 284 | 255.7433 | 329 | 250 | 230 | 124 | 255 | 266 | [7 empty] | 258.7861
| 1 | 5 | 261.6 | 286 | 239 | 257 | 258 | 276 | 249.1409 | 285 | 258.6221 | 255 | 240 | 264 | 213 | 251 | 256 | [7 empty] | 36
| 1 | 6 | 263.8 | 282 | 238 | 255 | 264 | 281 | 259.2593 | 284 | 262.1473 | 256 | 240 | 275 | 192 | 250 | 251
| Mean |  | 265.8667 | 276 | 237.1667 | 259.1667 | 262 | 276.8333 | 257.2408 | 285.1667 | 259.9124 | 272.5 | 243.3333 | 255.1667 | 187.8333 | 252.3333 | 258
| Median |  | 266.3 | 275.5 | 238 | 258.5 | 262 | 276 | 260.3635 | 285 | 260.3847 | 260 | 240 | 259 | 200 | 253 | 258
| Std Dev. |  | 2.9924 | 7.6681 | 2.3166 | 3.7639 | 3.0984 | 3.5449 | 6.6444 | 1.169 | 8.4884 | 28.6269 | 5.164 | 16.1792 | 34.0964 | 3.559 | 6.4187
| Rel.Std.Dev. |  | 1.126% | 2.778% | 0.9768% | 1.452% | 1.183% | 1.281% | 2.583% | 0.41% | 3.266% | 10.51% | 2.122% | 6.341% | 18.15% | 1.41% | 2.488%
| PDM3 |  | 2.736% | 6.652% | -8.354% | 0.147% | 1.242% | 6.974% | -0.5971% | 10.19% | 0.4352% | 5.299% | -5.971% | -1.399% | -27.42% | -2.493% | -0.3038%
| Z-Score (Absolute) |  | 0.66 | 1.66 | 2.17 | 0 | 0.28 | 1.74 | 0.19 | 2.57 | 0.07 | 1.32 | 1.56 | 0.39 | 7.04 | 0.67 | 0.12
| Analytical results for Be in OREAS 555b (Certified Value 5.33 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | [9 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | [9 empty] | 2
3
| 1 | 1 | 5.6 | 13 | 5 | 6 | < 5 | 5.5193 | 6 | 5.7235 | NA | 6 | 5 | 1 | 5 | [9 empty] | 1
| 1 | 2 | 5.2 | 13 | 5 | 5 | < 5 | 5.1822 | 6 | 5.4893 | NA | 5 | 5 | 2 | 5 | [9 empty] | 3
| 1 | 3 | 5.4 | 13 | 5 | 6 | < 5 | 5.1343 | 5 | 5.5778 | NA | 5 | 5 | 1 | 5 | [9 empty] | 16
| 1 | 4 | 5.4 | 12 | 5 | 6 | < 5 | 5.5276 | 5 | 5.3341 | NA | 5 | 5 | 1 | 5 | [9 empty] | 5.3275
| 1 | 5 | 5.5 | 11 | 6 | 6 | < 5 | 5.5581 | 6 | 5.3306 | NA | 5 | 5 | 2 | 5 | [9 empty] | 37
| 1 | 6 | 5.2 | 13 | 5 | 6 | < 5 | 5.59 | 5 | 5.4199 | NA | 5 | 5 | 1 | 5
| Mean |  | 5.3833 | 12.5 | 5.1667 | 5.8333 |  | 5.4186 | 5.5 | 5.4792 |  | 5.1667 | 5 | 1.3333 | 5
| Median |  | 5.4 | 13 | 5 | 6 |  | 5.5235 | 5.5 | 5.4546 |  | 5 | 5 | 1 | 5
| Std Dev. |  | 0.16021 | 0.83666 | 0.40825 | 0.40825 |  | 0.20378 | 0.54772 | 0.15241 |  | 0.40825 | 0 | 0.5164 | 0
| Rel.Std.Dev. |  | 2.976% | 6.693% | 7.902% | 6.999% |  | 3.761% | 9.959% | 2.782% |  | 7.902% | 0% | 38.73% | 0%
| PDM3 |  | 1.047% | 134.6% | -3.019% | 9.494% |  | 1.709% | 3.237% | 2.847% |  | -3.019% | -6.148% | -74.97% | -6.148%
| Z-Score (Absolute) |  | 0.29 | 19.49 | 0.29 | 1.51 | 7.48 | 0.39 | 0.61 | 0.55 | NA | 0.29 | 0.74 | 10.63 | 0.74
| Analytical results for Bi in OREAS 555b (Certified Value 1.56 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 15 | 17 | 19 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 5 | 1.6 | < 2 | 1.4 | 1.1327 | 1.72 | < 2 | 1.6 | 1.6 | [13 empty] | 1
| 1 | 2 | < 5 | 1.6 | < 2 | 1.5 | 1.0614 | 1.76 | 2 | 1.3 | 1.6 | [13 empty] | 4
| 1 | 3 | < 5 | 1.6 | < 2 | 1.6 | 1.045 | 1.71 | < 2 | 1.4 | 1.5 | [13 empty] | 16
| 1 | 4 | < 5 | 1.6 | < 2 | 1.5 | 1.0488 | 1.76 | < 2 | 1.4 | 1.5 | [13 empty] | 1.5558
| 1 | 5 | < 5 | 1.5 | < 2 | 1.5 | 1.1515 | 1.69 | < 2 | 1.4 | 1.6 | [13 empty] | 38
| 1 | 6 | < 5 | 1.6 | < 2 | 1.5 | 1.0329 | 1.73 | < 2 | 1.5 | 1.6
| Mean |  |  | 1.5833 |  | 1.5 | 1.0787 | 1.7283 | 2 | 1.4333 | 1.5667
| Median |  |  | 1.6 |  | 1.5 | 1.0551 | 1.725 | 2 | 1.4 | 1.6
| Std Dev. |  |  | 0.040825 |  | 0.063246 | 0.05029 | 0.027869 |  | 0.10328 | 0.05164
| Rel.Std.Dev. |  |  | 2.578% |  | 4.216% | 4.662% | 1.612% |  | 7.206% | 3.296%
| PDM3 |  |  | 1.772% |  | -3.585% | -30.66% | 11.09% | 28.55% | -7.87% | 0.7003%
| Z-Score (Absolute) |  | 2.95 | 0.25 | 1.48 | 0 | 1.24 | 0.67 | 0.98 | 0.2 | 0.2
| Analytical results for Ca in OREAS 555b (Certified Value 0.143 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.16 | 0.08 | 0.13 | 0.1 | 0.19 | 0.15 | 0.143 | 0.743 | 0.157 | 0.186 | 0.157 | 0.1276 | 0.6002 | 0.15 | 0.15846 | 0.2127 | 0.1 | 0.2 | 0.35 | 0.3 | 0.11 | 0.13 | 1
| 1 | 2 | 0.15 | 0.11 | 0.11 | 0.1 | 0.19 | 0.15 | 0.121 | 0.172 | 0.15 | 0.186 | 0.157 | 0.14227 | 0.6054 | 0.15 | 0.14512 | 0.1984 | 0.1 | 0.2 | 0.17 | 0.3 | 0.14 | 0.1 | #N/A
| 1 | 3 | 0.16 | 0.1 | 0.12 | 0.1 | 0.17 | 0.16 | 0.129 | 0.1 | 0.15 | 0.172 | 0.107 | 0.12239 | 0.6073 | 0.15 | 0.14241 | 0.1802 | 0.1 | 0.2 | 0.08 | 0.3 | 0.13 | 0.1 | 16
| 1 | 4 | 0.16 | 0.16 | 0.11 | 0.1 | 0.17 | 0.16 | 0.129 | 0.2 | 0.157 | 0.172 | 0.157 | 0.12907 | 0.6089 | 0.17 | 0.15277 | 0.3239 | 0.1 | 0.2 | 0.04 | 0.3 | 0.13 | 0.11 | 0.14299
| 1 | 5 | 0.15 | 0.16 | 0.12 | 0.1 | 0.16 | 0.15 | 0.121 | 0.143 | 0.143 | 0.172 | 0.157 | 0.12997 | 0.6087 | 0.16 | 0.14343 | 0.1692 | 0.1 | 0.2 | 0.11 | 0.3 | 0.13 | 0.1 | 39
| 1 | 6 | 0.17 | 0.14 | 0.12 | 0.1 | 0.19 | 0.17 | 0.107 | 0.172 | 0.143 | 0.172 | 0.157 | 0.13861 | 0.6015 | 0.16 | 0.14276 | 0.3213 | 0.1 | 0.2 | 0.08 | 0.3 | 0.14 | 0.13
| Mean |  | 0.15833 | 0.125 | 0.11833 | 0.1 | 0.17833 | 0.15667 | 0.125 | 0.255 | 0.15 | 0.17667 | 0.14867 | 0.13165 | 0.60533 | 0.15667 | 0.14749 | 0.23428 | 0.1 | 0.2 | 0.13833 | 0.3 | 0.13 | 0.11167
| Median |  | 0.16 | 0.125 | 0.12 | 0.1 | 0.18 | 0.155 | 0.125 | 0.172 | 0.15 | 0.172 | 0.157 | 0.12952 | 0.60635 | 0.155 | 0.14427 | 0.20555 | 0.1 | 0.2 | 0.095 | 0.3 | 0.13 | 0.105
| Std Dev. |  | 0.0075277 | 0.033317 | 0.0075277 | 1.5202e-17 | 0.013292 | 0.008165 | 0.011933 | 0.24146 | 0.006261 | 0.0072296 | 0.020412 | 0.007388 | 0.0037141 | 0.008165 | 0.0066117 | 0.070025 | 1.5202e-17 | 3.0405e-17 | 0.11232 | 0 | 0.010954 | 0.01472
| Rel.Std.Dev. |  | 4.754% | 26.65% | 6.361% | 1.52e-14% | 7.453% | 5.212% | 9.547% | 94.69% | 4.174% | 4.092% | 13.73% | 5.612% | 0.6136% | 5.212% | 4.483% | 29.89% | 1.52e-14% | 1.52e-14% | 81.2% | 0% | 8.427% | 13.18%
| PDM3 |  | 10.73% | -12.58% | -17.24% | -30.07% | 24.72% | 9.564% | -12.58% | 78.33% | 4.901% | 23.55% | 3.969% | -7.929% | 323.3% | 9.564% | 3.147% | 63.84% | -30.07% | 39.87% | -3.258% | 109.8% | -9.086% | -21.91%
| Z-Score (Absolute) |  | 0.27 | 0.75 | 0.96 | NA | 0.89 | 0.22 | 0.75 | 3.24 | 0.02 | 0.82 | 0.02 | 0.55 | 14 | 0.22 | 0.06 | 2.6 | NA | NA | 0.34 | NA | 0.6 | 1.16
Results from laboratory 21 were removed due to their 0.01 wt.% reading resolution. Results from laboratories 4, 23 and 26 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 555b (Indicative Value < 10 ppm) | [23 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 15 | 21 | 23 | 25 | 26 | [12 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 1 | < 5 | < 10 | < 2 | < 5 | < 10 | < 10 | 0.7 | 0.4 | < 10 | [12 empty] | 1
| 1 | 2 | < 1 | < 5 | < 10 | < 2 | < 5 | < 10 | < 10 | 0.7 | 0.2 | < 10 | [12 empty] | 5
| 1 | 3 | < 1 | < 5 | < 10 | < 2 | < 5 | < 10 | < 10 | 0.8 | 0.3 | < 10 | [12 empty] | 16
| 1 | 4 | < 1 | < 5 | < 10 | < 2 | < 5 | < 10 | < 10 | 0.7 | 0.3 | < 10 | [12 empty] | < 10
| 1 | 5 | < 1 | < 5 | < 10 | < 2 | < 5 | < 10 | < 10 | 0.7 | 0.4 | < 10 | [12 empty] | 24
| 1 | 6 | < 1 | < 5 | < 10 | < 2 | < 5 | < 10 | < 10 | 0.8 | 0.3 | < 10
| Mean | [8 empty] | 0.73333 | 0.31667
| Median | [8 empty] | 0.7 | 0.3
| Std Dev. | [8 empty] | 0.05164 | 0.075277
| Rel.Std.Dev. | [8 empty] | 7.042% | 23.77%
PDM3
| Z-Score (Absolute) |  | 0.64 | 0 | 0.81 | 0.48 | 0 | 0.81 | 0.81 | 0.57 | 0.7 | 0.81
| Analytical results for Ce in OREAS 555b (Certified Value 161 ppm) | [23 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 02 | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [14 empty] | 0
0
| 1 | 1 | 836 | 162 | 162 | 159.1783 | NA | 130 | 200 | 164.4 | [14 empty] | 1
| 1 | 2 | 905 | 160 | 164 | 167.5445 | NA | 147 | 184 | 158.4 | [14 empty] | 6
| 1 | 3 | 445 | 151 | 162 | 156.8649 | NA | 141 | 182 | 162.6 | [14 empty] | 16
| 1 | 4 | 452 | 159 | 156 | 161.4641 | NA | 134 | 156 | 166.4 | [14 empty] | 160.8185
| 1 | 5 | 1008 | 155 | 159 | 161.7279 | NA | 127 | 187 | 161.9 | [14 empty] | 40
| 1 | 6 | 846 | 161 | 163 | 165.787 | NA | 131 | 181 | 158.4
| Mean |  | 748.6667 | 158 | 161 | 162.0944 |  | 135 | 181.6667 | 162.0167
| Median |  | 841 | 159.5 | 162 | 161.596 |  | 132.5 | 183 | 162.25
| Std Dev. |  | 240.4243 | 4.1952 | 2.9665 | 3.993 |  | 7.5631 | 14.3481 | 3.2065
| Rel.Std.Dev. |  | 32.11% | 2.655% | 1.843% | 2.463% |  | 5.602% | 7.898% | 1.979%
| PDM3 |  | 365.5% | -1.753% | 0.1128% | 0.7934% |  | -16.05% | 12.96% | 0.745%
| Z-Score (Absolute) |  | 98.49 | 0.67 | 0.17 | 0.01 | NA | 4.54 | 3.3 | 0
| Analytical results for Co in OREAS 555b (Certified Value 1.11 wt.%) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 1.2389 | 1.1808 | 1.1234 | 1.1185 | 1.06 | 1.16 | 1.125 | 1.065 | 1.05 | 1.105 | 1.105 | 1.1495 | 1.0496 | 1.19 | 1.1362 | 1.1032 | 1.109 | 1.09 | 1.1403 | 1.056 | 1.126 |  | 1
| 1 | 2 | 1.2393 | 1.1361 | 1.127 | 1.1282 | 1.04 | 1.14 | 1.17 | 1.085 | 0.98 | 1.11 | 1.085 | 1.1438 | 1.0869 | 1.24 | 1.0912 | 1.1769 | 1.09 | 1.06 | 1.1366 | 1.052 | 1.144 |  | 36
| 1 | 3 | 1.1971 | 1.1547 | 1.1241 | 1.1022 | 1.07 | 1.13 | 1.08 | 1.065 | 1.06 | 1.12 | 1.075 | 1.121 | 1.0343 | 1.25 | 1.0922 | 1.0841 | 1.139 | 1.09 | 1.1462 | 1.083 | 1.171 |  | 16
| 1 | 4 | 1.196 | 1.1457 | 1.1238 | 1.1384 | 1.04 | 1.1 | 1.08 | 1.12 | 1.025 | 1.075 | 1.09 | 1.1272 | 1.038 | 1.26 | 1.0742 | 1.1224 | 1.13 | 1.06 | 1.1309 | 1.045 | 1.117 |  | 1.1053
| 1 | 5 | 1.2164 | 1.12 | 1.1184 | 1.0687 | 1.06 | 1.16 | 1.105 | 1.075 | 0.945 | 1.165 | 1.185 | 1.1316 | 1.0731 | 1.21 | 1.0695 | 1.1084 | 1.111 | 1.08 | 1.1334 | 1.076 | 1.124 |  | 41
| 1 | 6 | 1.1854 | 1.1646 | 1.1046 | 1.1193 | 1.06 | 1.08 | 1.115 | 1.15 | 0.988 | 1.12 | 1.11 | 1.1098 | 1.069 | 1.23 | 1.0891 | 1.1105 | 1.122 | 1.1 | 1.1119 | 1.041 | 1.155
| Mean |  | 1.2122 | 1.1503 | 1.1202 | 1.1125 | 1.055 | 1.1283 | 1.1125 | 1.0933 | 1.008 | 1.1158 | 1.1083 | 1.1305 | 1.0585 | 1.23 | 1.0921 | 1.1176 | 1.1168 | 1.08 | 1.1332 | 1.0588 | 1.1395
| Median |  | 1.2068 | 1.1502 | 1.1236 | 1.1189 | 1.06 | 1.135 | 1.11 | 1.08 | 1.0065 | 1.115 | 1.0975 | 1.1294 | 1.0593 | 1.235 | 1.0901 | 1.1095 | 1.1165 | 1.085 | 1.135 | 1.054 | 1.135
| Std Dev. |  | 0.023118 | 0.021424 | 0.0081394 | 0.024593 | 0.012247 | 0.032506 | 0.033578 | 0.034448 | 0.04452 | 0.029226 | 0.039707 | 0.014624 | 0.021048 | 0.026077 | 0.023603 | 0.031636 | 0.01736 | 0.016733 | 0.011747 | 0.016987 | 0.020888
| Rel.Std.Dev. |  | 1.907% | 1.862% | 0.7266% | 2.21% | 1.161% | 2.881% | 3.018% | 3.151% | 4.417% | 2.619% | 3.583% | 1.294% | 1.988% | 2.12% | 2.161% | 2.831% | 1.554% | 1.549% | 1.037% | 1.604% | 1.833%
| PDM3 |  | 9.665% | 4.069% | 1.345% | 0.6518% | -4.555% | 2.08% | 0.6473% | -1.087% | -8.807% | 0.9488% | 0.2703% | 2.274% | -4.24% | 11.28% | -1.201% | 1.107% | 1.039% | -2.293% | 2.521% | -4.208% | 3.09%
| Z-Score (Absolute) |  | 2.89 | 1.03 | 0.13 | 0.1 | 1.82 | 0.37 | 0.1 | 0.67 | 3.23 | 0 | 0.22 | 0.44 | 1.72 | 3.42 | 0.71 | 0.05 | 0.03 | 1.07 | 0.52 | 1.71 | 0.71
| Analytical results for Cr in OREAS 555b (Certified Value 112 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | 100 | 101 | 125 | 140 | 95 | 140 | 210 | 140 | 140 | 140 | 77.7544 | 110 | 118.0047 | 124 | 110 | 90 | 156 | 115 | 100 | < 100 |  |  | 1
| 1 | 2 | 100 | 102 | 121 | 140 | 97 | 140 | 70 | 140 | 210 | 140 | 81.1619 | 110 | 112.6554 | 131 | 110 | 125 | 165 | 115 | 100 | < 100 |  |  | 37
| 1 | 3 | 100 | 103 | 124 | 150 | 99 | 140 | < 70 | 140 | 140 | 70 | 82.0149 | 100 | 114.7329 | 124 | 100 | 125 | 167 | 118 | 100 | < 100 |  |  | 16
| 1 | 4 | 100 | 103 | 129 | 130 | 95 | 140 | 70 | 70 | 140 | 140 | 81.0199 | 100 | 113.3125 | 130 | 110 | 90 | 129 | 113 | 100 | < 100 |  |  | 111.7493
| 1 | 5 | 100 | 104 | 117 | 150 | 110 | 140 | 70 | 140 | 140 | 70 | 82.793 | 110 | 115.1981 | 130 | 110 | 119 | 167 | 114 | 100 | < 100 |  |  | 42
| 1 | 6 | 100 | 101 | 124 | 150 | 99 | 140 | 70 | 140 | 140 | 140 | 77.2919 | 110 | 111.9391 | 130 | 120 | 115 | 138 | 123 | 100 | < 100
| Mean |  | 100 | 102.3333 | 123.3333 | 143.3333 | 99.1667 | 140 | 98 | 128.3333 | 151.6667 | 116.6667 | 80.3393 | 106.6667 | 114.3071 | 128.1667 | 110 | 110.6667 | 153.6667 | 116.3333 | 100
| Median |  | 100 | 102.5 | 124 | 145 | 98 | 140 | 70 | 140 | 140 | 140 | 81.0909 | 110 | 114.0227 | 130 | 110 | 117 | 160.5 | 115 | 100
| Std Dev. |  | 0 | 1.2111 | 4.0332 | 8.165 | 5.6006 | 0 | 62.6099 | 28.5774 | 28.5774 | 36.1478 | 2.2776 | 5.164 | 2.1896 | 3.2506 | 6.3246 | 16.452 | 16.3911 | 3.6697 | 0
| Rel.Std.Dev. |  | 0% | 1.183% | 3.27% | 5.696% | 5.648% | 0% | 63.89% | 22.27% | 18.84% | 30.98% | 2.835% | 4.841% | 1.916% | 2.536% | 5.75% | 14.87% | 10.67% | 3.154% | 0%
| PDM3 |  | -10.51% | -8.426% | 10.37% | 28.26% | -11.26% | 25.28% | -12.3% | 14.84% | 35.72% | 4.4% | -28.11% | -4.548% | 2.289% | 14.69% | -1.565% | -0.9688% | 37.51% | 4.102% | -10.51%
| Z-Score (Absolute) |  | NA | 0.58 | 0.44 | NA | 0.74 | 1.1 | 1.4 | 0.54 | 1.65 | 0.01 | 1.65 | NA | 0 | 0.67 | NA | 0.18 | 1.91 | 0.1 | NA | 3.13
Results from laboratories 5, 17 and 21 were removed due to their 10 ppm reading resolution. Results from laboratories 2 and 27 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 555b (Certified Value 2.26 ppm) | [23 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | 28 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [11 empty] | 2
3
| 1 | 1 | 2.1 | 2.3 | 2 | 2.1936 | 2.2 | 2.419 | 2 | 2.1 | 2.8 | 2.2 | 2 | [11 empty] | 1
| 1 | 2 | 2.1 | 2.3 | 2 | 2.243 | 2.3 | 2.4899 | 2 | 2.4 | 2.6 | 2.1 | 2.2 | [11 empty] | 9
| 1 | 3 | 2 | 2.2 | 2.1 | 2.5761 | 2.2 | 2.2828 | 2 | 2.6 | 2.8 | 2.1 | 2.2 | [11 empty] | 16
| 1 | 4 | 2.3 | 2.1 | 2.1 | 2.3206 | 2.2 | 2.3409 | 2 | 2 | 2.4 | 2.1 | 2.2 | [11 empty] | 2.259
| 1 | 5 | 1.9 | 2.1 | 2.1 | 2.5115 | 2.2 | 2.316 | < 2 | 2.3 | 2.9 | 2 | 2.1 | [11 empty] | 43
| 1 | 6 | 2.1 | 2.4 | 2.2 | 2.4838 | 2.3 | 2.464 | < 2 | 2.2 | 2.6 | 2.1 | 2.1
| Mean |  | 2.0833 | 2.2333 | 2.0833 | 2.3881 | 2.2333 | 2.3854 | 2 | 2.2667 | 2.6833 | 2.1 | 2.1333
| Median |  | 2.1 | 2.25 | 2.1 | 2.4022 | 2.2 | 2.38 | 2 | 2.25 | 2.7 | 2.1 | 2.15
| Std Dev. |  | 0.13292 | 0.12111 | 0.075277 | 0.15695 | 0.05164 | 0.084314 | 0 | 0.21602 | 0.18348 | 0.063246 | 0.08165
| Rel.Std.Dev. |  | 6.38% | 5.423% | 3.613% | 6.572% | 2.312% | 3.535% | 0% | 9.531% | 6.838% | 3.012% | 3.827%
| PDM3 |  | -7.777% | -1.137% | -7.777% | 5.714% | -1.137% | 5.596% | -11.47% | 0.3386% | 18.78% | -7.039% | -5.564%
| Z-Score (Absolute) |  | 0.71 | 0 | 0.71 | 0.74 | 0 | 0.72 | NA | 0.16 | 2.14 | 0.63 | 0.48
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 555b (Certified Value 4.57 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 21 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 4.5202 | 4.439 | 4.7244 | 4.8625 | 4.26 | 4.62 | 4.62 | 4.49 | 4.58 | 4.6 | 4.72 | 4.4723 | 4.3425 | 4.72 | 4.3554 | 4.74 | 4.68 | 4.4982 | 4.521 | 4.705 |  |  | 1
| 1 | 2 | 4.5189 | 4.4044 | 4.7148 | 4.8185 | 4.16 | 4.61 | 4.59 | 4.58 | 4.41 | 4.56 | 4.66 | 4.4979 | 4.3514 | 4.69 | 4.5346 | 4.742 | 4.56 | 4.5837 | 4.414 | 4.766 |  |  | 39
| 1 | 3 | 4.5133 | 4.4124 | 4.7435 | 4.8215 | 4.34 | 4.64 | 4.64 | 4.52 | 4.5 | 4.49 | 4.69 | 4.5904 | 4.305 | 4.74 | 4.4112 | 4.712 | 4.62 | 4.472 | 4.422 | 4.632 |  |  | 16
| 1 | 4 | 4.5117 | 4.4182 | 4.7518 | 4.8469 | 4.31 | 4.63 | 4.68 | 4.56 | 4.62 | 4.45 | 4.68 | 4.4538 | 4.2924 | 4.68 | 4.548 | 4.679 | 4.35 | 4.5494 | 4.455 | 4.678 |  |  | 4.5665
| 1 | 5 | 4.521 | 4.4853 | 4.7467 | 4.648 | 4.58 | 4.64 | 4.66 | 4.44 | 4.38 | 4.5 | 4.69 | 4.4354 | 4.3193 | 4.8 | 4.5271 | 4.682 | 4.62 | 4.5096 | 4.406 | 4.592 |  |  | 44
| 1 | 6 | 4.5159 | 4.5095 | 4.709 | 4.7764 | 4.4 | 4.62 | 4.63 | 4.45 | 4.46 | 4.56 | 4.64 | 4.5814 | 4.3266 | 4.71 | 4.5121 | 4.756 | 4.73 | 4.4452 | 4.494 | 4.686
| Mean |  | 4.5168 | 4.4448 | 4.7317 | 4.7956 | 4.3417 | 4.6267 | 4.6367 | 4.5067 | 4.4917 | 4.5267 | 4.68 | 4.5052 | 4.3229 | 4.7233 | 4.4814 | 4.7185 | 4.5933 | 4.5097 | 4.452 | 4.6765
| Median |  | 4.5174 | 4.4286 | 4.734 | 4.82 | 4.325 | 4.625 | 4.635 | 4.505 | 4.48 | 4.53 | 4.685 | 4.4851 | 4.323 | 4.715 | 4.5196 | 4.726 | 4.62 | 4.5039 | 4.4385 | 4.682
| Std Dev. |  | 0.003828 | 0.04301 | 0.018014 | 0.078047 | 0.14204 | 0.012111 | 0.031411 | 0.057155 | 0.094322 | 0.055737 | 0.027568 | 0.065917 | 0.022232 | 0.043205 | 0.078868 | 0.032728 | 0.13261 | 0.050529 | 0.046891 | 0.060059
| Rel.Std.Dev. |  | 0.08475% | 0.9677% | 0.3807% | 1.627% | 3.272% | 0.2618% | 0.6775% | 1.268% | 2.1% | 1.231% | 0.5891% | 1.463% | 0.5143% | 0.9147% | 1.76% | 0.6936% | 2.887% | 1.12% | 1.053% | 1.284%
| PDM3 |  | -1.088% | -2.666% | 3.617% | 5.017% | -4.924% | 1.317% | 1.536% | -1.311% | -1.639% | -0.8728% | 2.485% | -1.343% | -5.336% | 3.434% | -1.864% | 3.328% | 0.5871% | -1.245% | -2.508% | 2.408%
| Z-Score (Absolute) |  | 0.04 | 0.57 | 1.56 | 2.03 | 1.34 | 0.78 | 0.85 | 0.11 | 0.22 | 0.04 | 1.17 | 0.12 | 1.47 | 1.49 | 0.3 | 1.46 | 0.53 | 0.09 | 0.52 | 1.15
| Analytical results for Dy in OREAS 555b (Certified Value 6.7 ppm) | [23 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 6.3 | 7.05 | 7.0219 | 9 | 6.24 | 7.64 | 6.85 | [15 empty] | 1
| 1 | 2 | 6.7 | 7.26 | 7.0079 | 8 | 6.8 | 6.05 | 6.67 | [15 empty] | 10
| 1 | 3 | 6.2 | 7.59 | 6.7352 | 9 | 6.49 | 6.53 | 7.04 | [15 empty] | 16
| 1 | 4 | 6.2 | 6.94 | 6.7477 | 8 | 6.65 | 5.55 | 7.08 | [15 empty] | 6.6976
| 1 | 5 | 6.4 | 7.1 | 6.5009 | 7 | 6.53 | 6.7 | 7.02 | [15 empty] | 45
| 1 | 6 | 6.3 | 7.33 | 7.07 | 7 | 6.4 | 5.97 | 6.87
| Mean |  | 6.35 | 7.2117 | 6.8473 | 8 | 6.5183 | 6.4067 | 6.9217
| Median |  | 6.3 | 7.18 | 6.8778 | 8 | 6.51 | 6.29 | 6.945
| Std Dev. |  | 0.18708 | 0.23319 | 0.22285 | 0.89443 | 0.19447 | 0.73132 | 0.15484
| Rel.Std.Dev. |  | 2.946% | 3.233% | 3.255% | 11.18% | 2.983% | 11.41% | 2.237%
| PDM3 |  | -5.19% | 7.675% | 2.235% | 19.45% | -2.677% | -4.344% | 3.345%
| Z-Score (Absolute) |  | 0.87 | 1.38 | 0.43 | NA | 0.43 | 0.72 | 0.63
| Analytical results for Er in OREAS 555b (Certified Value 3.83 ppm) | [23 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 3.6 | 4.1 | 3.839 | 5 | 3.44 | 4.05 | 3.98 | [15 empty] | 1
| 1 | 2 | 4 | 4.1 | 4.0147 | 4 | 3.82 | 3.69 | 3.86 | [15 empty] | 11
| 1 | 3 | 3.8 | 4.08 | 3.7689 | 5 | 3.54 | 4.2 | 4.02 | [15 empty] | 16
| 1 | 4 | 3.7 | 3.93 | 3.6211 | 4 | 3.54 | 3.45 | 4.09 | [15 empty] | 3.8268
| 1 | 5 | 3.7 | 4.06 | 3.6881 | 4 | 3.7 | 3.87 | 3.99 | [15 empty] | 46
| 1 | 6 | 3.9 | 4 | 3.822 | 4 | 3.53 | 3.45 | 4.09
| Mean |  | 3.7833 | 4.045 | 3.7923 | 4.3333 | 3.595 | 3.785 | 4.005
| Median |  | 3.75 | 4.07 | 3.7954 | 4 | 3.54 | 3.78 | 4.005
| Std Dev. |  | 0.1472 | 0.067454 | 0.13657 | 0.5164 | 0.13853 | 0.31085 | 0.085499
| Rel.Std.Dev. |  | 3.891% | 1.668% | 3.601% | 11.92% | 3.853% | 8.213% | 2.135%
| PDM3 |  | -1.135% | 5.703% | -0.9008% | 13.24% | -6.057% | -1.092% | 4.657%
| Z-Score (Absolute) |  | 0.04 | 1.74 | 0.02 | NA | 1.31 | 0.02 | 1.47
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 555b (Certified Value 0.97 ppm) | [23 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 1 | 0.96 | 1.0054 | < 2 | 0.89 | 1.04 | 1.05 | [15 empty] | 1
| 1 | 2 | 1 | 1 | 1.0415 | < 2 | 0.99 | 0.66 | 1.01 | [15 empty] | 12
| 1 | 3 | 0.9 | 1.02 | 1.0283 | < 2 | 0.93 | 0.85 | 1.04 | [15 empty] | 16
| 1 | 4 | 0.9 | 0.93 | 1.0232 | < 2 | 0.82 | 0.28 | 1.02 | [15 empty] | 0.9668
| 1 | 5 | 0.8 | 0.95 | 1.0215 | < 2 | 0.84 | 0.81 | 1.04 | [15 empty] | 47
| 1 | 6 | 0.9 | 0.94 | 1.0341 | < 2 | 0.92 | 0.55 | 1
| Mean |  | 0.91667 | 0.96667 | 1.0257 |  | 0.89833 | 0.69833 | 1.0267
| Median |  | 0.9 | 0.955 | 1.0257 |  | 0.905 | 0.735 | 1.03
| Std Dev. |  | 0.075277 | 0.03559 | 0.012366 |  | 0.062423 | 0.26483 | 0.019664
| Rel.Std.Dev. |  | 8.212% | 3.682% | 1.206% |  | 6.949% | 37.92% | 1.915%
| PDM3 |  | -5.185% | -0.01377% | 6.089% |  | -7.082% | -27.77% | 6.192%
| Z-Score (Absolute) |  | 0.57 | 0 | 0.67 | 0.38 | 0.78 | 3.07 | 0.69
| Analytical results for Fe in OREAS 555b (Certified Value 3.87 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 3.78 | 3.93 | 3.72 | 4.04 | 3.85 | 4 | 3.819 | 3.749 | 3.917 | 3.833 | 3.959 | 4.079 | 4.1612 | 4.4 | 3.9676 | 3.7687 | 3.87 | 3.89 | 3.64 | 3.7 | 3.72 | 3.98 | 1
| 1 | 2 | 3.7 | 3.84 | 3.73 | 4.03 | 3.85 | 3.95 | 3.861 | 3.847 | 3.728 | 3.833 | 3.903 | 4.1164 | 4.1055 | 4.49 | 3.9022 | 3.9418 | 3.8 | 4.1 | 3.7 | 3.78 | 3.68 | 3.97 | #N/A
| 1 | 3 | 3.61 | 3.86 | 3.74 | 4.01 | 3.85 | 4.01 | 3.812 | 3.749 | 3.763 | 3.833 | 3.889 | 3.9821 | 4.1068 | 4.29 | 3.7857 | 3.8323 | 3.85 | 4.08 | 3.8 | 3.69 | 3.64 | 3.87 | 16
| 1 | 4 | 3.74 | 3.84 | 3.72 | 4.02 | 3.86 | 4.06 | 3.812 | 3.861 | 3.847 | 3.707 | 3.868 | 4.1356 | 4.1992 | 4.33 | 3.8624 | 3.9282 | 3.86 | 4.16 | 3.53 | 3.75 | 3.7 | 3.92 | 3.8742
| 1 | 5 | 3.71 | 3.77 | 3.74 | 3.84 | 3.85 | 4.05 | 3.812 | 3.777 | 3.77 | 3.861 | 3.889 | 4.12 | 4.1866 | 4.48 | 3.8837 | 3.9182 | 3.78 | 4.14 | 3.69 | 3.71 | 3.67 | 3.91 | 48
| 1 | 6 | 3.7 | 3.84 | 3.75 | 3.98 | 3.82 | 4.07 | 3.833 | 3.721 | 3.882 | 3.882 | 3.91 | 3.9782 | 4.1903 | 4.36 | 3.8477 | 3.9074 | 3.87 | 4.16 | 3.94 | 3.67 | 3.7 | 3.85
| Mean |  | 3.7067 | 3.8467 | 3.7333 | 3.9867 | 3.8467 | 4.0233 | 3.8248 | 3.784 | 3.8178 | 3.8248 | 3.903 | 4.0685 | 4.1583 | 4.3917 | 3.8749 | 3.8828 | 3.8383 | 4.0883 | 3.7167 | 3.7167 | 3.685 | 3.9167
| Median |  | 3.705 | 3.84 | 3.735 | 4.015 | 3.85 | 4.03 | 3.8155 | 3.763 | 3.8085 | 3.833 | 3.896 | 4.0977 | 4.1739 | 4.38 | 3.8731 | 3.9128 | 3.855 | 4.12 | 3.695 | 3.705 | 3.69 | 3.915
| Std Dev. |  | 0.056451 | 0.051251 | 0.012111 | 0.074744 | 0.013663 | 0.045461 | 0.019508 | 0.057212 | 0.075056 | 0.061091 | 0.03099 | 0.07095 | 0.042299 | 0.08085 | 0.060458 | 0.067778 | 0.038687 | 0.10245 | 0.14038 | 0.040825 | 0.028107 | 0.052026
| Rel.Std.Dev. |  | 1.523% | 1.332% | 0.3244% | 1.875% | 0.3552% | 1.13% | 0.51% | 1.512% | 1.966% | 1.597% | 0.794% | 1.744% | 1.017% | 1.841% | 1.56% | 1.746% | 1.008% | 2.506% | 3.777% | 1.098% | 0.7627% | 1.328%
| PDM3 |  | -4.326% | -0.7119% | -3.637% | 2.902% | -0.7119% | 3.848% | -1.275% | -2.329% | -1.456% | -1.275% | 0.7421% | 5.015% | 7.331% | 13.36% | 0.01652% | 0.2199% | -0.927% | 5.526% | -4.067% | -4.067% | -4.885% | 1.095%
| Z-Score (Absolute) |  | 1.03 | 0 | 0.83 | 1.03 | 0 | 1.3 | 0.16 | 0.46 | 0.21 | 0.16 | 0.41 | 1.63 | 2.29 | 4.01 | 0.21 | 0.27 | 0.06 | 1.78 | 0.96 | 0.96 | 1.19 | 0.51
| Analytical results for Ga in OREAS 555b (Certified Value 29.1 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [10 empty] | 1
2
| 1 | 1 | 35 | 30 | 26.2 | 30.5 | 29.2041 | 30.1 | 29.9926 | 26 | 25 | 36 | 28 | 21 | [10 empty] | 1
| 1 | 2 | 35 | 29 | 27.5 | 30.6 | 30.0445 | 30.4 | 29.1384 | 28 | 26 | 31 | 29 | 23 | [10 empty] | 13
| 1 | 3 | 35.7 | 29 | 27.5 | 31.1 | 28.9273 | 30.7 | 29.7484 | 27 | 25 | 28 | 31 | 22 | [10 empty] | 16
| 1 | 4 | 35.1 | 29 | 28.9 | 28.9 | 30.2058 | 30.1 | 30.0148 | 28 | 26 | 19 | 29 | 21 | [10 empty] | 29.1167
| 1 | 5 | 34.6 | 29 | 27.6 | 30 | 30.0767 | 30 | 29.4351 | 27 | 26 | 32 | 30 | 22 | [10 empty] | 49
| 1 | 6 | 35.6 | 30 | 28 | 29.5 | 29.6322 | 30 | 29.2835 | 27 | 25 | 25 | 29 | 22
| Mean |  | 35.1667 | 29.3333 | 27.6167 | 30.1 | 29.6818 | 30.2167 | 29.6021 | 27.1667 | 25.5 | 28.5 | 29.3333 | 21.8333
| Median |  | 35.05 | 29 | 27.55 | 30.25 | 29.8383 | 30.1 | 29.5917 | 27 | 25.5 | 29.5 | 29 | 22
| Std Dev. |  | 0.41312 | 0.5164 | 0.87502 | 0.8025 | 0.52192 | 0.27869 | 0.37113 | 0.75277 | 0.54772 | 5.9582 | 1.0328 | 0.75277
| Rel.Std.Dev. |  | 1.175% | 1.76% | 3.168% | 2.666% | 1.758% | 0.9223% | 1.254% | 2.771% | 2.148% | 20.91% | 3.521% | 3.448%
| PDM3 |  | 20.78% | 0.7439% | -5.152% | 3.377% | 1.941% | 3.778% | 1.667% | -6.697% | -12.42% | -2.118% | 0.7439% | -25.01%
| Z-Score (Absolute) |  | 4.58 | 0 | 1.35 | 0.6 | 0.27 | 0.69 | 0.21 | 1.7 | 3.01 | 0.65 | 0 | 5.89
| Analytical results for Gd in OREAS 555b (Certified Value 7.26 ppm) | [23 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
2
| 1 | 1 | 7.8 | 7.27 | 7.755 | 9 | 6.36 | 8.6 | 7 | [15 empty] | 1
| 1 | 2 | 7.4 | 7.21 | 7.8482 | 8 | 7.54 | 7.74 | 7 | [15 empty] | 14
| 1 | 3 | 6.7 | 7.36 | 7.5337 | 9 | 6.74 | 8.55 | 7 | [15 empty] | 16
| 1 | 4 | 7.3 | 7.11 | 7.3846 | 8 | 6.39 | 5.58 | 7 | [15 empty] | 7.257
| 1 | 5 | 6.8 | 7.15 | 7.3384 | 7 | 6.5 | 9.06 | 8 | [15 empty] | 50
| 1 | 6 | 7 | 7.08 | 7.77 | 8 | 6.46 | 7.43 | 8
| Mean |  | 7.1667 | 7.1967 | 7.605 | 8.1667 | 6.665 | 7.8267 | 7.3333
| Median |  | 7.15 | 7.18 | 7.6443 | 8 | 6.48 | 8.145 | 7
| Std Dev. |  | 0.41312 | 0.10539 | 0.21616 | 0.75277 | 0.44925 | 1.2534 | 0.5164
| Rel.Std.Dev. |  | 5.764% | 1.464% | 2.842% | 9.218% | 6.741% | 16.01% | 7.042%
| PDM3 |  | -1.245% | -0.8312% | 4.795% | 12.54% | -8.158% | 7.85% | 1.052%
| Z-Score (Absolute) |  | 0.05 | 0 | 0.67 | NA | 0.88 | 1.04 | NA
Results from laboratories 19 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 555b (Certified Value 2.01 ppm) | [23 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [12 empty] | 2
3
| 1 | 1 | 2 | 2.1 | 3 | 2.0449 | 2 | 1.963 | NA | 2 | < 5 | 2 | [12 empty] | 1
| 1 | 2 | 2 | 2.2 | 3 | 2.1631 | 2.1 | 2.0078 | NA | 2 | < 5 | 2 | [12 empty] | 15
| 1 | 3 | 2 | 2.5 | 3 | 1.8608 | 2 | 2.0232 | NA | 2 | < 5 | 2 | [12 empty] | 16
| 1 | 4 | 2 | 2.7 | 2 | 2.0515 | 2 | 1.9876 | NA | 2 | < 5 | 2 | [12 empty] | 2.0081
| 1 | 5 | 2 | 2.5 | 3 | 2.4391 | 1.9 | 2.076 | NA | 2 | < 5 | 2 | [12 empty] | 51
| 1 | 6 | 2 | 2.3 | 3 | 2.539 | 2 | 2.0548 | NA | 2 | < 5 | 2
| Mean |  | 2 | 2.3833 | 2.8333 | 2.183 | 2 | 2.0187 |  | 2 |  | 2
| Median |  | 2 | 2.4 | 3 | 2.1073 | 2 | 2.0155 |  | 2 |  | 2
| Std Dev. |  | 0 | 0.22286 | 0.40825 | 0.25804 | 0.063246 | 0.04193 |  | 0 |  | 0
| Rel.Std.Dev. |  | 0% | 9.351% | 14.41% | 11.82% | 3.162% | 2.077% |  | 0% |  | 0%
| PDM3 |  | -0.4049% | 18.68% | 41.09% | 8.71% | -0.4049% | 0.5284% |  | -0.4049% |  | -0.4049%
| Z-Score (Absolute) |  | 0.67 | 13.12 | 29.31 | 5.91 | 0.67 | 0 | NA | 0.67 | 17.31 | 0.67
| Analytical results for Hf in OREAS 555b (Indicative Value 6.84 ppm) | [23 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 25 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | [17 empty] | 2
2
| 1 | 1 | 10 | 6 | 7.0249 | 6 | 11 | [17 empty] | 1
| 1 | 2 | 10 | 7 | 6.9415 | 7 | 9 | [17 empty] | 19
| 1 | 3 | < 10 | 8 | 7.0397 | 7 | 10 | [17 empty] | 16
| 1 | 4 | < 10 | 6 | 7.0384 | 7 | 8 | [17 empty] | 6.8398
| 1 | 5 | < 10 | 7 | 7.0857 | 7 | 10 | [17 empty] | 25
| 1 | 6 | < 10 | 7 | 7.2734 | 7 | 9
| Mean |  | 10 | 6.8333 | 7.0673 | 6.8333 | 9.5
| Median |  | 10 | 7 | 7.0391 | 7 | 9.5
| Std Dev. |  | 0 | 0.75277 | 0.11139 | 0.40825 | 1.0488
| Rel.Std.Dev. |  | 0% | 11.02% | 1.576% | 5.974% | 11.04%
| PDM3 |  | 46.2% | -0.09521% | 3.325% | -0.09521% | 38.89%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.95 | 0 | 10.79
| Analytical results for Hg in OREAS 555b (Indicative Value < 0.1 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [20 empty] | 1
| No | No | 01 | 25 | [20 empty] | ppm
|  |  | PF*OES | PF*OES | [20 empty] | 2
2
| 1 | 1 | 5.11 | < 0.1 | [20 empty] | 1
| 1 | 2 | < 5 | < 0.1 | [20 empty] | 20
| 1 | 3 | < 5 | < 0.1 | [20 empty] | 16
| 1 | 4 | 5.83 | < 0.1 | [20 empty] | < 0.1
| 1 | 5 | < 5 | < 0.1 | [20 empty] | 26
| 1 | 6 | 6.48 | < 0.1
| Mean |  | 5.8067
| Median |  | 5.83
| Std Dev. |  | 0.6853
| Rel.Std.Dev. |  | 11.8%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 555b (Certified Value 1.38 ppm) | [23 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 1.4 | 1.38 | 1.469 | < 2 | 1.15 | 1.61 | 1.44 | [15 empty] | 1
| 1 | 2 | 1.4 | 1.42 | 1.3398 | < 2 | 1.3 | 1.41 | 1.37 | [15 empty] | 3
| 1 | 3 | 1.2 | 1.41 | 1.2223 | < 2 | 1.23 | 1.35 | 1.4 | [15 empty] | 16
| 1 | 4 | 1.3 | 1.36 | 1.3228 | < 2 | 1.16 | 1.13 | 1.41 | [15 empty] | 1.3779
| 1 | 5 | 1.3 | 1.37 | 1.2875 | < 2 | 1.25 | 1.63 | 1.42 | [15 empty] | 52
| 1 | 6 | 1.4 | 1.39 | 1.4455 | < 2 | 1.23 | 1.34 | 1.41
| Mean |  | 1.3333 | 1.3883 | 1.3478 |  | 1.22 | 1.4117 | 1.4083
| Median |  | 1.35 | 1.385 | 1.3313 |  | 1.23 | 1.38 | 1.41
| Std Dev. |  | 0.08165 | 0.023166 | 0.094122 |  | 0.056569 | 0.18723 | 0.023166
| Rel.Std.Dev. |  | 6.124% | 1.669% | 6.983% |  | 4.637% | 13.26% | 1.645%
| PDM3 |  | -3.234% | 0.7575% | -2.183% |  | -11.46% | 2.451% | 2.209%
| Z-Score (Absolute) |  | 0.16 | 0.45 | 0 | 3.88 | 1.42 | 0.71 | 0.67
| Analytical results for In in OREAS 555b (Certified Value 2.47 ppm) | [23 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 17 | 19 | 23 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 2.5 | 2.4 | 2.4 | 2.6972 | 2.58 | 2.4 | 2.2 | 3 | 2.6 | [13 empty] | 1
| 1 | 2 | 2.5 | 2.4 | 2.6 | 2.5668 | 2.59 | 2.6 | 2.4 | 2.7 | 2.5 | [13 empty] | 4
| 1 | 3 | 2.4 | 2.5 | 2.5 | 2.3505 | 2.64 | 2.5 | 2.2 | 2.9 | 2.4 | [13 empty] | 16
| 1 | 4 | 2.5 | 2.3 | 2.5 | 2.7258 | 2.6 | 2.5 | 2.1 | 2.7 | 2.4 | [13 empty] | 2.4674
| 1 | 5 | 2.3 | 2.1 | 2.6 | 2.6846 | 2.57 | 2.5 | 2.3 | 2.8 | 2.6 | [13 empty] | 53
| 1 | 6 | 2.6 | 2.5 | 2.4 | 2.418 | 2.61 | 2.6 | 2.2 | 2.8 | 2.4
| Mean |  | 2.4667 | 2.3667 | 2.5 | 2.5738 | 2.5983 | 2.5167 | 2.2333 | 2.8167 | 2.4833
| Median |  | 2.5 | 2.4 | 2.5 | 2.6257 | 2.595 | 2.5 | 2.2 | 2.8 | 2.45
| Std Dev. |  | 0.10328 | 0.15055 | 0.089443 | 0.15799 | 0.024833 | 0.075277 | 0.10328 | 0.1169 | 0.098319
| Rel.Std.Dev. |  | 4.187% | 6.361% | 3.578% | 6.138% | 0.9557% | 2.991% | 4.624% | 4.15% | 3.959%
| PDM3 |  | -0.02775% | -4.081% | 1.323% | 4.315% | 5.309% | 1.999% | -9.485% | 14.16% | 0.6477%
| Z-Score (Absolute) |  | 0.3 | 1.22 | 0 | 0.67 | 0.9 | 0.15 | 2.44 | 2.89 | 0.15
| Analytical results for K in OREAS 555b (Certified Value 2.58 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 2.47 | 2.92 | 2.54 | 2.68 | 2.6 | 2.4 | 2.607 | 2.665 | 2.424 | 2.623 | 2.656 | 2.6925 | 2.9547 | 3.02 | 2.4053 | 2.4 | 3.1 | 2.68 | 2.69 | 2.55 | 2.6 |  | 1
| 1 | 2 | 2.44 | 2.92 | 2.53 | 2.76 | 2.7 | 2.39 | 2.59 | 2.582 | 2.316 | 2.573 | 2.623 | 2.6892 | 2.9204 | 3.16 | 2.5079 | 2.4 | 3.1 | 2.69 | 2.74 | 2.53 | 2.6 |  | #N/A
| 1 | 3 | 2.36 | 3.01 | 2.55 | 2.66 | 2.6 | 2.41 | 2.598 | 2.507 | 2.358 | 2.549 | 2.607 | 2.599 | 2.9027 | 2.99 | 2.4333 | 2.4 | 3.1 | 2.92 | 2.71 | 2.47 | 2.6 |  | 16
| 1 | 4 | 2.45 | 3.15 | 2.56 | 2.7 | 2.6 | 2.41 | 2.607 | 2.64 | 2.507 | 2.524 | 2.632 | 2.7123 | 2.9127 | 3.09 | 2.4828 | 2.4 | 2.8 | 2.98 | 2.73 | 2.56 | 2.6 |  | 2.5809
| 1 | 5 | 2.43 | 3.32 | 2.56 | 2.57 | 2.6 | 2.39 | 2.607 | 2.499 | 2.3 | 2.549 | 2.632 | 2.5899 | 2.9471 | 3.15 | 2.4366 | 2.5 | 3.1 | 2.5 | 2.69 | 2.48 | 2.6 |  | 54
| 1 | 6 | 2.42 | 3.17 | 2.53 | 2.65 | 2.6 | 2.41 | 2.598 | 2.549 | 2.391 | 2.582 | 2.623 | 2.5782 | 2.9218 | 3.09 | 2.4698 | 2.4 | 3.1 | 3.08 | 2.69 | 2.53 | 2.7
| Mean |  | 2.4283 | 3.0817 | 2.545 | 2.67 | 2.6167 | 2.4017 | 2.6012 | 2.5737 | 2.3827 | 2.5667 | 2.6288 | 2.6435 | 2.9266 | 3.0833 | 2.456 | 2.4167 | 3.05 | 2.8083 | 2.7083 | 2.52 | 2.6167
| Median |  | 2.435 | 3.08 | 2.545 | 2.67 | 2.6 | 2.405 | 2.6025 | 2.5655 | 2.3745 | 2.561 | 2.6275 | 2.6441 | 2.9211 | 3.09 | 2.4532 | 2.4 | 3.1 | 2.805 | 2.7 | 2.53 | 2.6
| Std Dev. |  | 0.037639 | 0.15917 | 0.013784 | 0.06261 | 0.040825 | 0.0098319 | 0.0070261 | 0.068468 | 0.07632 | 0.034332 | 0.016142 | 0.060555 | 0.020178 | 0.06802 | 0.037588 | 0.040825 | 0.12247 | 0.21967 | 0.022286 | 0.036878 | 0.040825
| Rel.Std.Dev. |  | 1.55% | 5.165% | 0.5416% | 2.345% | 1.56% | 0.4094% | 0.2701% | 2.66% | 3.203% | 1.338% | 0.614% | 2.291% | 0.6895% | 2.206% | 1.531% | 1.689% | 4.016% | 7.822% | 0.8229% | 1.463% | 1.56%
| PDM3 |  | -5.911% | 19.4% | -1.39% | 3.453% | 1.387% | -6.944% | 0.786% | -0.2796% | -7.68% | -0.5508% | 1.858% | 2.427% | 13.39% | 19.47% | -4.841% | -6.363% | 18.18% | 8.813% | 4.938% | -2.359% | 1.387%
| Z-Score (Absolute) |  | 1.31 | 3.24 | 0.5 | 0.37 | 0 | 1.5 | 0.11 | 0.3 | 1.63 | 0.35 | 0.08 | 0.19 | 2.16 | 3.26 | 1.12 | 1.4 | 3.02 | 1.34 | 0.64 | 0.67 | 0
| Analytical results for La in OREAS 555b (Certified Value 77 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 18 | 19 | 21 | 23 | 25 | 28 | [11 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | [11 empty] | 0
1
| 1 | 1 | 78 | 66 | 74 | 81.8 | 81.1 | 85.4816 | 130 | 80 | 73.6 | 80.9 | 84.64 | [11 empty] | 1
| 1 | 2 | 77.4 | 69 | 73.4 | 81.3 | 83.2 | 84.7865 | 88 | 70 | 79.2 | 72.8 | 81.14 | [11 empty] | 5
| 1 | 3 | 78.3 | 61 | 74 | 79.5 | 81.6 | 76.7084 | 127 | 80 | 76.3 | 76.6 | 82.99 | [11 empty] | 16
| 1 | 4 | 77.9 | 58 | 73 | 80 | 79.1 | 79.5244 | 98 | 80 | 72.7 | 61 | 85.26 | [11 empty] | 76.5578
| 1 | 5 | 76.7 | 67 | 75.8 | 77.4 | 79.8 | 78.2376 | 76 | 70 | 73 | 78.7 | 83.33 | [11 empty] | 55
| 1 | 6 | 77.4 | 64 | 73.8 | 80.1 | 82.1 | 84.9244 | 83 | 80 | 73.1 | 71.7 | 81.26
| Mean |  | 77.6167 | 64.1667 | 74 | 80.0167 | 81.15 | 81.6105 | 100.3333 | 76.6667 | 74.65 | 73.6167 | 83.1033
| Median |  | 77.65 | 65 | 73.9 | 80.05 | 81.35 | 82.1555 | 93 | 80 | 73.35 | 74.7 | 83.16
| Std Dev. |  | 0.57067 | 4.0702 | 0.96333 | 1.5459 | 1.5057 | 3.8939 | 22.9841 | 5.164 | 2.5867 | 7.0895 | 1.6929
| Rel.Std.Dev. |  | 0.7352% | 6.343% | 1.302% | 1.932% | 1.855% | 4.771% | 22.91% | 6.736% | 3.465% | 9.63% | 2.037%
| PDM3 |  | 1.383% | -16.19% | -3.341% | 4.518% | 5.998% | 6.6% | 31.06% | 0.1422% | -2.492% | -3.842% | 8.55%
| Z-Score (Absolute) |  | 0.19 | 2.34 | 0.77 | 0.19 | 0.37 | 0.45 | 3.43 | NA | 0.66 | 0.83 | 0.68
Results from laboratory 21 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 555b (Certified Value 60 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 15 | 17 | 19 | 21 | 25 | 26 | 27 | 28 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | [9 empty] | 0
1
| 1 | 1 | 65 | 54 | 55 | 63 | 62 | 59.1404 | 62.1 | 55 | 80 | 53.4 | 59 | 60 | < 100 | [9 empty] | 1
| 1 | 2 | 62 | 53.4 | 56 | 65 | 62 | 63.0293 | 63.7 | 57 | 80 | 54.5 | 59 | 60 | < 100 | [9 empty] | 33
| 1 | 3 | 66 | 53.2 | 56 | 64 | 62 | 66.7576 | 62.8 | 68 | 70 | 55 | 59 | 60 | < 100 | [9 empty] | 16
| 1 | 4 | 64 | 53 | 56 | 64 | 62 | 61.2964 | 61 | 64 | 70 | 54.2 | 59 | 60 | < 100 | [9 empty] | 59.7113
| 1 | 5 | 68 | 54 | 55 | 64 | 62 | 65.0679 | 62.4 | 55 | 70 | 50 | 61 | 50 | < 100 | [9 empty] | 56
| 1 | 6 | 66 | 54 | 56 | 64 | 62 | 59.4687 | 63.8 | 55 | 80 | 46.7 | 58 | 60 | < 100
| Mean |  | 65.1667 | 53.6 | 55.6667 | 64 | 62 | 62.46 | 62.6333 | 59 | 75 | 52.3 | 59.1667 | 58.3333
| Median |  | 65.5 | 53.7 | 56 | 64 | 62 | 62.1629 | 62.6 | 56 | 75 | 53.8 | 59 | 60
| Std Dev. |  | 2.0412 | 0.45607 | 0.5164 | 0.63246 | 0 | 3.0632 | 1.052 | 5.6214 | 5.4772 | 3.2741 | 0.98319 | 4.0825
| Rel.Std.Dev. |  | 3.132% | 0.8509% | 0.9277% | 0.9882% | 0% | 4.904% | 1.68% | 9.528% | 7.303% | 6.26% | 1.662% | 6.999%
| PDM3 |  | 9.136% | -10.23% | -6.774% | 7.182% | 3.833% | 4.603% | 4.894% | -1.191% | 25.6% | -12.41% | -0.9122% | -2.308%
| Z-Score (Absolute) |  | 1.16 | 1.07 | 0.67 | 0.93 | 0.55 | 0.63 | 0.67 | 0.03 | NA | 1.32 | 0 | NA | 1.77
Results from laboratories 21 and 27 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 555b (Certified Value 0.58 ppm) | [23 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 06 | 18 | 19 | 23 | 25 | 28 | [16 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [16 empty] | 2
2
| 1 | 1 | 0.49 | 0.55233 | NA | 0.74 | 0.58 | 0.57 | [16 empty] | 1
| 1 | 2 | 0.52 | 0.56423 | NA | 0.78 | 0.54 | 0.54 | [16 empty] | 25
| 1 | 3 | 0.5 | 0.52788 | NA | 0.72 | 0.52 | 0.57 | [16 empty] | 16
| 1 | 4 | 0.48 | 0.52074 | NA | 0.73 | 0.47 | 0.59 | [16 empty] | 0.57565
| 1 | 5 | 0.48 | 0.53869 | NA | 0.78 | 0.57 | 0.55 | [16 empty] | 57
| 1 | 6 | 0.48 | 0.55563 | NA | 0.73 | 0.53 | 0.55
| Mean |  | 0.49167 | 0.54325 |  | 0.74667 | 0.535 | 0.56167
| Median |  | 0.485 | 0.54551 |  | 0.735 | 0.535 | 0.56
| Std Dev. |  | 0.016021 | 0.016969 |  | 0.026583 | 0.03937 | 0.018348
| Rel.Std.Dev. |  | 3.258% | 3.124% |  | 3.56% | 7.359% | 3.267%
| PDM3 |  | -14.59% | -5.628% |  | 29.71% | -7.062% | -2.429%
| Z-Score (Absolute) |  | 1.89 | 0 | NA | 7.45 | 0.3 | 0.67
| Analytical results for Mg in OREAS 555b (Certified Value 3.09 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 3.14 | 3.14 | 2.93 | 3.22 | 2.99 | 3.18 | 3.027 | 3.064 | 2.919 | 3.142 | 3.16 | 3.2849 | 4.9585 | 3.14 | 3.2508 | 3.0941 | 3.11 | 3.11 | 3.38 | 3.01 | 3.01 | 3.21 | 1
| 1 | 2 | 3.1 | 3 | 2.9 | 3.23 | 3.06 | 3.16 | 3.027 | 3.088 | 2.828 | 3.082 | 3.148 | 3.271 | 4.9141 | 3.2 | 3.168 | 3.1692 | 3.06 | 3.24 | 3.26 | 3.07 | 2.98 | 3.2 | #N/A
| 1 | 3 | 3.02 | 3.02 | 2.93 | 3.2 | 3.02 | 3.18 | 3.009 | 3.058 | 2.889 | 3.094 | 3.118 | 3.177 | 4.9797 | 3.16 | 3.1036 | 3.1023 | 3.04 | 3.03 | 3.34 | 3 | 2.94 | 3.11 | 16
| 1 | 4 | 3.13 | 3.02 | 2.9 | 3.22 | 3.03 | 3.18 | 3.039 | 3.07 | 2.937 | 3.027 | 3.142 | 3.3172 | 4.9179 | 3.28 | 3.1491 | 3.1419 | 3.12 | 2.91 | 3.25 | 3.05 | 2.98 | 3.13 | 3.0941
| 1 | 5 | 3.09 | 2.95 | 2.92 | 3.07 | 3.04 | 3.2 | 3.027 | 3.009 | 2.84 | 3.058 | 3.172 | 3.1864 | 4.9415 | 3.19 | 3.0718 | 3.1504 | 3.05 | 3.3 | 3.23 | 3.02 | 2.93 | 3.16 | 58
| 1 | 6 | 3.07 | 2.97 | 2.94 | 3.19 | 3.03 | 3.17 | 3.039 | 2.997 | 2.889 | 3.076 | 3.142 | 3.1779 | 4.9616 | 3.17 | 3.0595 | 3.1453 | 2.93 | 3.29 | 3.47 | 2.99 | 2.98 | 3.18
| Mean |  | 3.0917 | 3.0167 | 2.92 | 3.1883 | 3.0283 | 3.1783 | 3.028 | 3.0477 | 2.8837 | 3.0798 | 3.147 | 3.2357 | 4.9455 | 3.19 | 3.1338 | 3.1339 | 3.0517 | 3.1467 | 3.3217 | 3.0233 | 2.97 | 3.165
| Median |  | 3.095 | 3.01 | 2.925 | 3.21 | 3.03 | 3.18 | 3.027 | 3.061 | 2.889 | 3.079 | 3.145 | 3.2287 | 4.95 | 3.18 | 3.1263 | 3.1436 | 3.055 | 3.175 | 3.3 | 3.015 | 2.98 | 3.17
| Std Dev. |  | 0.043551 | 0.066533 | 0.016733 | 0.059805 | 0.023166 | 0.013292 | 0.011009 | 0.036225 | 0.042791 | 0.038348 | 0.018363 | 0.062467 | 0.025929 | 0.04899 | 0.07124 | 0.029308 | 0.067946 | 0.15731 | 0.092826 | 0.030768 | 0.029665 | 0.03937
| Rel.Std.Dev. |  | 1.409% | 2.206% | 0.5731% | 1.876% | 0.765% | 0.4182% | 0.3636% | 1.189% | 1.484% | 1.245% | 0.5835% | 1.931% | 0.5243% | 1.536% | 2.273% | 0.9352% | 2.227% | 4.999% | 2.795% | 1.018% | 0.9988% | 1.244%
| PDM3 |  | -0.07729% | -2.501% | -5.626% | 3.047% | -2.124% | 2.724% | -2.135% | -1.499% | -6.8% | -0.4597% | 1.711% | 4.579% | 59.84% | 3.101% | 1.285% | 1.287% | -1.37% | 1.7% | 7.356% | -2.286% | -4.01% | 2.293%
| Z-Score (Absolute) |  | 0.19 | 0.85 | 1.7 | 0.67 | 0.74 | 0.58 | 0.74 | 0.58 | 2.02 | 0.29 | 0.3 | 1.09 | 16.17 | 0.68 | 0.19 | 0.19 | 0.54 | 0.3 | 1.84 | 0.79 | 1.26 | 0.46
| Analytical results for Mn in OREAS 555b (Certified Value 0.093 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.08702 | 0.1 | 0.0884 | 0.0957 | 0.0971 | 0.0927 | 0.093 | 0.093 | 0.093 | 0.093 | 0.093 | 0.096236 | 0.0978 | 0.091 | 0.095569 | 0.0953 | 0.086 | 0.0923 | 0.085 | 0.0875 | 0.09 | 0.09 | 1
| 1 | 2 | 0.08425 | 0.09 | 0.0886 | 0.0999 | 0.0953 | 0.0918 | 0.101 | 0.093 | 0.085 | 0.093 | 0.093 | 0.095279 | 0.0975 | 0.094 | 0.089674 | 0.1052 | 0.094 | 0.0946 | 0.088 | 0.0888 | 0.09 | 0.09 | #N/A
| 1 | 3 | 0.07866 | 0.09 | 0.0882 | 0.0966 | 0.0917 | 0.0906 | 0.093 | 0.093 | 0.093 | 0.093 | 0.093 | 0.091693 | 0.0972 | 0.092 | 0.094536 | 0.0926 | 0.09 | 0.1092 | 0.092 | 0.0896 | 0.09 | 0.09 | 16
| 1 | 4 | 0.08229 | 0.09 | 0.0882 | 0.1002 | 0.0947 | 0.0951 | 0.093 | 0.101 | 0.085 | 0.093 | 0.093 | 0.09163 | 0.0975 | 0.094 | 0.092199 | 0.0958 | 0.089 | 0.0835 | 0.093 | 0.0908 | 0.09 | 0.09 | 0.0925
| 1 | 5 | 0.08271 | 0.09 | 0.0885 | 0.0922 | 0.0969 | 0.102 | 0.093 | 0.093 | 0.077 | 0.093 | 0.108 | 0.095388 | 0.0974 | 0.093 | 0.087412 | 0.0938 | 0.088 | 0.1005 | 0.09 | 0.0873 | 0.09 | 0.09 | 59
| 1 | 6 | 0.0814 | 0.1 | 0.0885 | 0.0987 | 0.0934 | 0.0913 | 0.093 | 0.101 | 0.085 | 0.093 | 0.093 | 0.092099 | 0.097 | 0.093 | 0.090608 | 0.0951 | 0.089 | 0.0979 | 0.097 | 0.0866 | 0.09 | 0.09
| Mean |  | 0.082722 | 0.093333 | 0.0884 | 0.097217 | 0.09485 | 0.093917 | 0.094333 | 0.095667 | 0.086333 | 0.093 | 0.0955 | 0.093721 | 0.0974 | 0.092833 | 0.091667 | 0.0963 | 0.089333 | 0.096333 | 0.090833 | 0.088433 | 0.09 | 0.09
| Median |  | 0.0825 | 0.09 | 0.08845 | 0.09765 | 0.095 | 0.09225 | 0.093 | 0.093 | 0.085 | 0.093 | 0.093 | 0.093689 | 0.09745 | 0.093 | 0.091404 | 0.0952 | 0.089 | 0.09625 | 0.091 | 0.08815 | 0.09 | 0.09
| Std Dev. |  | 0.0028015 | 0.005164 | 0.00016733 | 0.003038 | 0.0020744 | 0.0042555 | 0.003266 | 0.0041312 | 0.0060222 | 1.5202e-17 | 0.0061237 | 0.0021286 | 0.00027568 | 0.001169 | 0.0030638 | 0.0045131 | 0.0026583 | 0.0085946 | 0.0041673 | 0.0015883 | 1.5202e-17 | 1.5202e-17
| Rel.Std.Dev. |  | 3.387% | 5.533% | 0.1893% | 3.125% | 2.187% | 4.531% | 3.462% | 4.318% | 6.975% | 1.635e-14% | 6.412% | 2.271% | 0.283% | 1.259% | 3.342% | 4.686% | 2.976% | 8.922% | 4.588% | 1.796% | 1.689e-14% | 1.689e-14%
| PDM3 |  | -10.57% | 0.9009% | -4.432% | 5.099% | 2.541% | 1.532% | 1.982% | 3.423% | -6.667% | 0.5405% | 3.243% | 1.32% | 5.297% | 0.3603% | -0.9011% | 4.108% | -3.423% | 4.144% | -1.802% | -4.396% | -2.703% | -2.703%
| Z-Score (Absolute) |  | 2.55 | 0.05 | 1.16 | 1 | 0.42 | 0.19 | 0.27 | 0.58 | 1.63 | 0.05 | 0.58 | 0.14 | 1.04 | 0.07 | 0.36 | 0.77 | 0.93 | 0.78 | 0.56 | 1.15 | 0.77 | 0.77
| Analytical results for Mo in OREAS 555b (Certified Value 17.3 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | [9 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | [9 empty] | 1
1
| 1 | 1 | 16.7 | 15.9 | 17 | 19 | 17 | 19.83 | 18 | 17.1611 | 22 | 20 | 17 | 14 | 18 | [9 empty] | 1
| 1 | 2 | 17 | 15.3 | 18 | 18 | 15 | 18.6888 | 19 | 17.533 | 20 | 20 | 17 | 13 | 18 | [9 empty] | 7
| 1 | 3 | 17 | 15.8 | 16 | 19 | 15 | 18.2547 | 18 | 16.577 | 20 | 10 | 17 | 13 | 18 | [9 empty] | 16
| 1 | 4 | 17.4 | 15.9 | 17 | 18 | 15 | 19.5398 | 18 | 17.0184 | 48 | 20 | 16 | 12 | 18 | [9 empty] | 17.2971
| 1 | 5 | 16.5 | 15.9 | 18 | 19 | 15 | 18.6274 | 19 | 16.8016 | 23 | 10 | 17 | 14 | 20 | [9 empty] | 60
| 1 | 6 | 16.7 | 15.7 | 17 | 20 | 15 | 18.2623 | 18 | 16.96 | 19 | 20 | 17 | 14 | 18
| Mean |  | 16.8833 | 15.75 | 17.1667 | 18.8333 | 15.3333 | 18.8672 | 18.3333 | 17.0085 | 25.3333 | 16.6667 | 16.8333 | 13.3333 | 18.3333
| Median |  | 16.85 | 15.85 | 17 | 19 | 15 | 18.6581 | 18 | 16.9892 | 21 | 20 | 17 | 13.5 | 18
| Std Dev. |  | 0.31885 | 0.23452 | 0.75277 | 0.75277 | 0.8165 | 0.6648 | 0.5164 | 0.32553 | 11.2012 | 5.164 | 0.40825 | 0.8165 | 0.8165
| Rel.Std.Dev. |  | 1.889% | 1.489% | 4.385% | 3.997% | 5.325% | 3.524% | 2.817% | 1.914% | 44.22% | 30.98% | 2.425% | 6.124% | 4.454%
| PDM3 |  | -2.392% | -8.944% | -0.7543% | 8.881% | -11.35% | 9.077% | 5.991% | -1.669% | 46.46% | -3.645% | -2.681% | -22.92% | 5.991%
| Z-Score (Absolute) |  | 0.11 | 0.7 | 0.04 | 0.91 | 0.92 | 0.93 | 0.65 | 0.04 | 4.3 | NA | 0.13 | 1.96 | 0.65
Results from laboratory 21 were removed due to their 10 ppm reading resolution.
| Analytical results for Na in OREAS 555b (Indicative Value 0.05 wt.%) | [23 empty] | Indicative
Na | Round | Replicate | Lab | [21 empty] | 1
| No | No | 01 | [21 empty] | wt.%
|  |  | PF*OES | [21 empty] | 3
3
| 1 | 1 | 0.05 | [21 empty] | 1
| 1 | 2 | 0.05 | [21 empty] | 17
| 1 | 3 | 0.05 | [21 empty] | 16
| 1 | 4 | 0.05 | [21 empty] | 0.05
| 1 | 5 | 0.05 | [21 empty] | 23
| 1 | 6 | 0.05
| Mean |  | 0.05
| Median |  | 0.05
| Std Dev. |  | 7.6012e-18
| Rel.Std.Dev. |  | 1.52e-14%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 555b (Certified Value 25.4 ppm) | [23 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 02 | 05 | 06 | 15 | 18 | 19 | 25 | 28 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | [14 empty] | 1
2
| 1 | 1 | 25 | 29 | 25 | 23.5284 | 26.2326 | 24 | 26.2 | < 20 | [14 empty] | 1
| 1 | 2 | 25 | 27.4 | 25 | 27.2208 | 25.423 | 26 | 26 | < 20 | [14 empty] | 9
| 1 | 3 | 26 | 25.7 | 26 | 26.3904 | 25.161 | 25 | 27.7 | < 20 | [14 empty] | 16
| 1 | 4 | 26 | 29 | 24 | 28.3267 | 25.1196 | 26 | 25.6 | < 20 | [14 empty] | 25.3883
| 1 | 5 | 25 | 28 | 25 | 23.9907 | 25.084 | 25 | 29.2 | < 20 | [14 empty] | 61
| 1 | 6 | 20 | 28.9 | 25 | 26.7557 | 25.0164 | 25 | 27.4 | < 20
| Mean |  | 24.5 | 28 | 25 | 26.0354 | 25.3394 | 25.1667 | 27.0167
| Median |  | 25 | 28.45 | 25 | 26.573 | 25.1403 | 25 | 26.8
| Std Dev. |  | 2.2583 | 1.3008 | 0.63246 | 1.885 | 0.45924 | 0.75277 | 1.3482
| Rel.Std.Dev. |  | 9.218% | 4.646% | 2.53% | 7.24% | 1.812% | 2.991% | 4.99%
| PDM3 |  | -3.499% | 10.29% | -1.529% | 2.549% | -0.1925% | -0.873% | 6.414%
| Z-Score (Absolute) |  | 0.66 | 2.41 | 0.22 | 0.69 | 0.08 | 0.08 | 1.55 | 13.4
| Analytical results for Nd in OREAS 555b (Certified Value 68 ppm) | [23 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 0
1
| 1 | 1 | 74.2 | 64 | 67.0985 | 68 | 53.2 | 78.9 | 69.1 | [15 empty] | 1
| 1 | 2 | 74 | 65 | 65.21 | 67 | 56.1 | 69.8 | 67.5 | [15 empty] | 10
| 1 | 3 | 69.6 | 65 | 61.3377 | 70 | 56.7 | 72.5 | 68.2 | [15 empty] | 16
| 1 | 4 | 69.3 | 63 | 61.4132 | 68 | 57.2 | 58.7 | 69.6 | [15 empty] | 67.539
| 1 | 5 | 69.8 | 63 | 63.4341 | 62 | 53.3 | 75.1 | 69.3 | [15 empty] | 62
| 1 | 6 | 71.4 | 65 | 63.609 | 65 | 53.7 | 67.2 | 70.1
| Mean |  | 71.3833 | 64.1667 | 63.6837 | 66.6667 | 55.0333 | 70.3667 | 68.9667
| Median |  | 70.6 | 64.5 | 63.5216 | 67.5 | 54.9 | 71.15 | 69.2
| Std Dev. |  | 2.2275 | 0.98319 | 2.2234 | 2.8048 | 1.8305 | 7.0176 | 0.95429
| Rel.Std.Dev. |  | 3.12% | 1.532% | 3.491% | 4.207% | 3.326% | 9.973% | 1.384%
| PDM3 |  | 5.692% | -4.993% | -5.708% | -1.292% | -18.52% | 4.187% | 2.114%
| Z-Score (Absolute) |  | 1.07 | 0.57 | 0.67 | 0 | 2.63 | 0.84 | 0.52
| Analytical results for Ni in OREAS 555b (Certified Value 228 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 0
0
| 1 | 1 | 207.6 | 242 | 220 | 219 | 260 | 234 | 220 | 290 | 240 | 240 | 200 | 267.7746 | 307 | 230 | 210.1363 | NA | 210 | 209 | 231 | 224 | 260 | 260 | 1
| 1 | 2 | 202.9 | 241 | 222 | 230 | 250 | 235 | 250 | 210 | 200 | 230 | 210 | 268.6366 | 305 | 240 | 211.4306 | NA | 200 | 225 | 205 | 214 | 230 | 260 | 40
| 1 | 3 | 192.5 | 253 | 221 | 220 | 260 | 237 | 220 | 190 | 250 | 250 | 200 | 263.767 | 302 | 240 | 214.0689 | NA | 200 | 218 | 213 | 217 | 230 | 260 | 16
| 1 | 4 | 197.5 | 248 | 220 | 234 | 250 | 240 | 230 | 260 | 210 | 240 | 210 | 269.3271 | 310 | 240 | 206.4511 | NA | 210 | 209 | 206 | 223 | 250 | 270 | 228.1801
| 1 | 5 | 199.9 | 231 | 223 | 208 | 250 | 338 | 220 | 210 | 210 | 250 | 240 | 261.3404 | 307 | 240 | 208.6059 | NA | 200 | 226 | 261 | 211 | 250 | 270 | 63
| 1 | 6 | 204.5 | 236 | 223 | 222 | 240 | 225 | 220 | 230 | 230 | 240 | 210 | 263.88 | 301 | 250 | 211.6887 | NA | 180 | 219 | 212 | 208 | 220 | 260
| Mean |  | 200.8167 | 241.8333 | 221.5 | 222.1667 | 251.6667 | 251.5 | 226.6667 | 231.6667 | 223.3333 | 241.6667 | 211.6667 | 265.7876 | 305.3333 | 240 | 210.3969 |  | 200 | 217.6667 | 221.3333 | 216.1667 | 240 | 263.3333
| Median |  | 201.4 | 241.5 | 221.5 | 221 | 250 | 236 | 220 | 220 | 220 | 240 | 210 | 265.8273 | 306 | 240 | 210.7835 |  | 200 | 218.5 | 212.5 | 215.5 | 240 | 260
| Std Dev. |  | 5.3838 | 7.9352 | 1.3784 | 9.1305 | 7.5277 | 42.6743 | 12.1106 | 37.1035 | 19.6638 | 7.5277 | 14.7196 | 3.2278 | 3.3862 | 6.3246 | 2.6482 |  | 10.9545 | 7.4207 | 21.5654 | 6.4317 | 15.4919 | 5.164
| Rel.Std.Dev. |  | 2.681% | 3.281% | 0.6223% | 4.11% | 2.991% | 16.97% | 5.343% | 16.02% | 8.805% | 3.115% | 6.954% | 1.214% | 1.109% | 2.635% | 1.259% |  | 5.477% | 3.409% | 9.743% | 2.975% | 6.455% | 1.961%
| PDM3 |  | -11.99% | 5.984% | -2.928% | -2.635% | 10.29% | 10.22% | -0.6632% | 1.528% | -2.124% | 5.911% | -7.237% | 16.48% | 33.81% | 5.18% | -7.793% |  | -12.35% | -4.607% | -3.001% | -5.265% | 5.18% | 15.41%
| Z-Score (Absolute) |  | 1.16 | 0.68 | 0.23 | 0.2 | 1.12 | 1.12 | 0 | 0.22 | 0.15 | 0.67 | 0.67 | 1.76 | 3.54 | 0.6 | 0.73 | NA | 1.2 | 0.4 | 0.24 | 0.47 | 0.6 | 1.65
| Analytical results for P in OREAS 555b (Certified Value 0.04 wt.%) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 03 | 06 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | [11 empty] | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [11 empty] | 3
3
| 1 | 1 | 0.04 | 0.04 | 0.0785 | 0.05 | 0.04 | < 0.05 | 0.0355 | 0.03 | 0.1 | 0.06 | 0.04 | [11 empty] | 1
| 1 | 2 | 0.04 | 0.04 | 0.0786 | 0.05 | 0.04 | < 0.05 | 0.0327 | 0.03 | 0.11 | 0.06 | 0.03 | [11 empty] | #N/A
| 1 | 3 | 0.04 | 0.04 | 0.0782 | 0.05 | 0.04 | < 0.05 | 0.0336 | 0.03 | 0.1 | 0.06 | 0.04 | [11 empty] | 16
| 1 | 4 | 0.03 | 0.04 | 0.0785 | 0.05 | 0.04 | < 0.05 | 0.0342 | 0.04 | 0.09 | 0.03 | 0.04 | [11 empty] | 0.040115
| 1 | 5 | 0.04 | 0.04 | 0.0786 | 0.06 | 0.04 | < 0.05 | 0.0345 | 0.03 | 0.1 | 0.05 | 0.04 | [11 empty] | 64
| 1 | 6 | 0.04 | 0.03 | 0.0779 | 0.05 | 0.04 | < 0.05 | 0.035 | 0.03 | 0.1 | 0.03 | 0.04
| Mean |  | 0.038333 | 0.038333 | 0.078383 | 0.051667 | 0.04 |  | 0.03425 | 0.031667 | 0.1 | 0.048333 | 0.038333
| Median |  | 0.04 | 0.04 | 0.0785 | 0.05 | 0.04 |  | 0.03435 | 0.03 | 0.1 | 0.055 | 0.04
| Std Dev. |  | 0.0040825 | 0.0040825 | 0.00027869 | 0.0040825 | 0 |  | 0.0010015 | 0.0040825 | 0.0063246 | 0.01472 | 0.0040825
| Rel.Std.Dev. |  | 10.65% | 10.65% | 0.3555% | 7.902% | 0% |  | 2.924% | 12.89% | 6.325% | 30.45% | 10.65%
| PDM3 |  | -4.44% | -4.44% | 95.4% | 28.8% | -0.2856% |  | -14.62% | -21.06% | 149.3% | 20.49% | -4.44%
| Z-Score (Absolute) |  | 0 | 0 | 4.05 | 1.35 | 0.17 | 1.35 | 0.41 | 0.67 | 6.24 | 1.01 | 0
| Analytical results for Pb in OREAS 555b (Certified Value < 100 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 21 | 23 | 25 | 26 | 27 | 28 | [5 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 0
0
| 1 | 1 | 15.7 | < 100 | 19.3 | 14 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | 18 | 20 | 26 | 30 | < 50 | < 300 | < 100 | [5 empty] | 1
| 1 | 2 | 14.6 | < 100 | 22.6 | 14 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 18 | 20 | 27 | 15 | < 50 | < 300 | < 100 | [5 empty] | 11
| 1 | 3 | 14 | < 100 | 20.2 | 14 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 19 | 20 | 25 | 10 | < 50 | < 300 | < 100 | [5 empty] | 16
| 1 | 4 | 14.4 | < 100 | 18.3 | 14 | < 100 | 1000 | < 100 | < 100 | 400 | < 20 | 17 | 20 | 27 | 23 | < 50 | < 300 | < 100 | [5 empty] | < 100
| 1 | 5 | 14.9 | < 100 | 18.7 | 14 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 18 | 20 | 26 | 18 | < 50 | < 300 | < 100 | [5 empty] | 65
| 1 | 6 | 16.8 | < 100 | 18.9 | 14 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 17 | 20 | 27 | 11 | < 50 | < 300 | < 100
| Mean |  | 15.0667 |  | 19.6667 | 14 |  | 1000 |  |  | 250 |  | 17.8333 | 20 | 26.3333 | 17.8333
| Median |  | 14.75 |  | 19.1 | 14 |  | 1000 |  |  | 250 |  | 18 | 20 | 26.5 | 16.5
| Std Dev. |  | 1.0231 |  | 1.5756 | 0 |  |  |  |  | 212.132 |  | 0.75277 | 0 | 0.8165 | 7.6267
| Rel.Std.Dev. |  | 6.79% |  | 8.012% | 0% |  |  |  |  | 84.85% |  | 4.221% | 0% | 3.101% | 42.77%
PDM3
| Z-Score (Absolute) |  | 0.47 | 0.98 | 0.28 | 0.51 | 0.98 | 7.51 | 0.98 | 0.98 | 3.73 | 0.67 | 0.35 | 0.26 | 0 | 0.35 | 0.06 | 5.11 | 0.98
| Analytical results for Pr in OREAS 555b (Certified Value 18.6 ppm) | [23 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 1
2
| 1 | 1 | 18.7 | 16.8 | 19.2006 | 19 | 15.76 | 20.9 | 18.7 | [15 empty] | 1
| 1 | 2 | 18.9 | 16.8 | 18.8053 | 19 | 17.73 | 18.8 | 17.88 | [15 empty] | 12
| 1 | 3 | 17.9 | 17.1 | 18.1712 | 19 | 16.64 | 19.9 | 18.51 | [15 empty] | 16
| 1 | 4 | 18.8 | 16.2 | 18.1736 | 19 | 15.05 | 16.3 | 18.75 | [15 empty] | 18.6327
| 1 | 5 | 18.3 | 16.2 | 17.6671 | 17 | 15.57 | 20.6 | 18.6 | [15 empty] | 66
| 1 | 6 | 18.9 | 17 | 18.4926 | 18 | 16.01 | 18.8 | 18.23
| Mean |  | 18.5833 | 16.6833 | 18.4184 | 18.5 | 16.1267 | 19.2167 | 18.445
| Median |  | 18.75 | 16.8 | 18.3331 | 19 | 15.885 | 19.35 | 18.555
| Std Dev. |  | 0.40208 | 0.392 | 0.53898 | 0.83666 | 0.94354 | 1.6774 | 0.33219
| Rel.Std.Dev. |  | 2.164% | 2.35% | 2.926% | 4.522% | 5.851% | 8.729% | 1.801%
| PDM3 |  | -0.2648% | -10.46% | -1.15% | -0.7121% | -13.45% | 3.134% | -1.007%
| Z-Score (Absolute) |  | 0.67 | 8.59 | 0.13 | 0.27 | 11.3 | 3.76 | 0
| Analytical results for Rb in OREAS 555b (Certified Value 96 ppm) | [23 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | 28 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [11 empty] | 0
1
| 1 | 1 | 95.9 | 97.4 | 92 | 92.238 | 101 | 99.5977 | 92 | 85 | 0.103 | 103.1 | 98 | [11 empty] | 1
| 1 | 2 | 95.4 | 98.9 | 93 | 92.3811 | 106 | 96.7565 | 95 | 88 | 0.096 | 101.5 | 93 | [11 empty] | 13
| 1 | 3 | 92.7 | 96.9 | 94 | 91.5596 | 104 | 96.892 | 96 | 85 | 0.102 | 99.9 | 91 | [11 empty] | 16
| 1 | 4 | 94.9 | 97.9 | 93 | 93.7979 | 107 | 97.2674 | 101 | 85 | 0.093 | 100.3 | 95 | [11 empty] | 95.6344
| 1 | 5 | 91.9 | 92.6 | 91 | 93.7615 | 102 | 96.2136 | 98 | 87 | 0.102 | 102.8 | 92 | [11 empty] | 67
| 1 | 6 | 95.4 | 99 | 93 | 92.6805 | 106 | 97 | 98 | 86 | 0.098 | 97 | 92
| Mean |  | 94.3667 | 97.1167 | 92.6667 | 92.7364 | 104.3333 | 97.2879 | 96.6667 | 86 | 0.099 | 100.7667 | 93.5
| Median |  | 95.15 | 97.65 | 93 | 92.5308 | 105 | 96.946 | 97 | 85.5 | 0.1 | 100.9 | 92.5
| Std Dev. |  | 1.6513 | 2.3609 | 1.0328 | 0.88771 | 2.4221 | 1.1842 | 3.0768 | 1.2649 | 0.004 | 2.2483 | 2.5884
| Rel.Std.Dev. |  | 1.75% | 2.431% | 1.115% | 0.9572% | 2.322% | 1.217% | 3.183% | 1.471% | 4.04% | 2.231% | 2.768%
| PDM3 |  | -1.326% | 1.55% | -3.103% | -3.03% | 9.096% | 1.729% | 1.079% | -10.07% | -99.9% | 5.367% | -2.232%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.42 | 0.4 | 2.44 | 0.72 | 0.56 | 2.05 | 23.11 | 1.57 | 0.21
| Analytical results for Re in OREAS 555b (Indicative Value 0.059 ppm) | [23 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 3
3
| 1 | 1 | < 0.1 | < 0.1 | < 0.05 | < 0.01 | 0.05 | 0.07 | 34.2 | < 0.1 | [14 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.05 | < 0.01 | 0.04 | 0.07 | 40.4 | < 0.1 | [14 empty] | 14
| 1 | 3 | < 0.1 | < 0.1 | < 0.05 | < 0.01 | 0.03 | 0.07 | 33.5 | < 0.1 | [14 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.05 | < 0.01 | 0.08 | 0.08 | 31.8 | < 0.1 | [14 empty] | 0.059167
| 1 | 5 | < 0.1 | < 0.1 | < 0.05 | < 0.01 | 0.04 | 0.06 | 25.5 | < 0.1 | [14 empty] | 24
| 1 | 6 | < 0.1 | < 0.1 | < 0.05 | < 0.01 | 0.05 | 0.07 | 34 | < 0.1
| Mean | [5 empty] | 0.048333 | 0.07 | 33.2333
| Median | [5 empty] | 0.045 | 0.07 | 33.75
| Std Dev. | [5 empty] | 0.017224 | 0.0063246 | 4.7919
| Rel.Std.Dev. | [5 empty] | 35.64% | 9.035% | 14.42%
| PDM3 | [5 empty] | -18.31% | 18.31% | 56069%
| Z-Score (Absolute) |  | 0 | 0 | 1.56 | 2.8 | 0.1 | 1.24 | 2065.46 | 0
| Analytical results for S in OREAS 555b (Certified Value 4.45 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 3.78 | 4 | 4.53 | 4.55 | 4.34 | 4.44 | 4.51 | 4.48 | 4.36 | 4.55 | 4.56 | 4.8737 | 5.612 | 4.55 | 4.2325 | 4.38 | 4.46 | 4.73 | 4.38 | 4.24 | 4.57 |  | 1
| 1 | 2 | 3.6 | 4 | 4.55 | 4.47 | 4.45 | 4.42 | 4.51 | 4.42 | 4.23 | 4.47 | 4.53 | 4.8794 | 5.6522 | 4.66 | 4.3328 | 4.36 | 4.81 | 4.16 | 4.45 | 4.25 | 4.49 |  | 28
| 1 | 3 | 3.47 | 4 | 4.56 | 4.52 | 4.33 | 4.44 | 4.52 | 4.38 | 4.32 | 4.5 | 4.54 | 4.7142 | 5.6083 | 4.56 | 4.121 | 4.43 | 4.5 | 4.39 | 4.38 | 3.94 | 4.29 |  | 16
| 1 | 4 | 3.44 | 4 | 4.53 | 4.45 | 4.54 | 4.46 | 4.39 | 4.41 | 4.48 | 4.48 | 4.57 | 4.6149 | 5.6111 | 4.76 | 4.2399 | 4.41 | 4.27 | 4.12 | 4.42 | 4.29 | 4.43 |  | 4.4524
| 1 | 5 | 3.64 | 4 | 4.54 | 4.34 | 4.34 | 4.49 | 4.45 | 4.32 | 4.3 | 4.43 | 4.52 | 4.539 | 5.6372 | 4.5 | 4.3466 | 4.41 | 4.5 | 4.35 | 4.36 | 4.21 | 4.42 |  | 68
| 1 | 6 | 3.21 | 4 | 4.55 | 4.51 | 4.46 | 4.43 | 4.48 | 4.36 | 4.4 | 4.51 | 4.58 | 4.6554 | 5.6303 | 4.5 | 4.2934 | 4.45 | 4.84 | 4.54 | 4.38 | 4.26 | 4.51
| Mean |  | 3.5233 | 4 | 4.5433 | 4.4733 | 4.41 | 4.4467 | 4.4767 | 4.395 | 4.3483 | 4.49 | 4.55 | 4.7128 | 5.6252 | 4.5883 | 4.261 | 4.4067 | 4.5633 | 4.3817 | 4.395 | 4.1983 | 4.4517
| Median |  | 3.535 | 4 | 4.545 | 4.49 | 4.395 | 4.44 | 4.495 | 4.395 | 4.34 | 4.49 | 4.55 | 4.6848 | 5.6212 | 4.555 | 4.2667 | 4.41 | 4.5 | 4.37 | 4.38 | 4.245 | 4.46
| Std Dev. |  | 0.19664 | 0 | 0.012111 | 0.074476 | 0.086255 | 0.025033 | 0.049666 | 0.055045 | 0.086352 | 0.040497 | 0.023664 | 0.13909 | 0.01765 | 0.10245 | 0.082933 | 0.03266 | 0.22006 | 0.23025 | 0.033317 | 0.12922 | 0.096419
| Rel.Std.Dev. |  | 5.581% | 0% | 0.2666% | 1.665% | 1.956% | 0.563% | 1.109% | 1.252% | 1.986% | 0.9019% | 0.5201% | 2.951% | 0.3138% | 2.233% | 1.946% | 0.7411% | 4.822% | 5.255% | 0.7581% | 3.078% | 2.166%
| PDM3 |  | -20.87% | -10.16% | 2.042% | 0.4694% | -0.9531% | -0.1295% | 0.5443% | -1.29% | -2.338% | 0.8437% | 2.191% | 5.847% | 26.34% | 3.052% | -4.299% | -1.028% | 2.491% | -1.589% | -1.29% | -5.707% | -0.01723%
| Z-Score (Absolute) |  | 6.44 | 3.12 | 0.67 | 0.19 | 0.26 | 0 | 0.21 | 0.36 | 0.69 | 0.3 | 0.72 | 1.86 | 8.22 | 0.99 | 1.29 | 0.28 | 0.81 | 0.45 | 0.36 | 1.73 | 0.03
| Analytical results for Sb in OREAS 555b (Certified Value 4 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 15 | 21 | 23 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | < 5 | 4.2 | 4 | 4 | 3.6403 | < 50 | 3 | 4 | 4.2 | [13 empty] | 1
| 1 | 2 | < 5 | 3.4 | 4 | 4 | 3.8726 | < 50 | 3 | 4 | 3.9 | [13 empty] | 15
| 1 | 3 | < 5 | 3.5 | 4 | 4 | 4.3978 | < 50 | 3 | 4 | 4 | [13 empty] | 16
| 1 | 4 | < 5 | 4.2 | 4 | 4 | 4.1899 | < 50 | 3 | 4 | 4.2 | [13 empty] | 4.0001
| 1 | 5 | < 5 | 4 | 4 | 4 | 3.6471 | < 50 | 3 | 4 | 4 | [13 empty] | 69
| 1 | 6 | < 5 | 3.8 | 4 | 4 | 4.0541 | < 50 | 3 | 4 | 3.9
| Mean |  |  | 3.85 | 4 | 4 | 3.967 |  | 3 | 4 | 4.0333
| Median |  |  | 3.9 | 4 | 4 | 3.9634 |  | 3 | 4 | 4
| Std Dev. |  |  | 0.34496 | 0 | 0 | 0.3036 |  | 0 | 0 | 0.13663
| Rel.Std.Dev. |  |  | 8.96% | 0% | 0% | 7.653% |  | 0% | 0% | 3.387%
| PDM3 |  |  | -3.751% | -0.001478% | -0.001478% | -0.8274% |  | -25% | -0.001478% | 0.8318%
| Z-Score (Absolute) |  | 30.34 | 3.03 | 0 | 0 | 0.67 | 424.81 | 20.23 | 0 | 0.67
| Analytical results for Sc in OREAS 555b (Certified Value 11.8 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 04 | 06 | 15 | 17 | 18 | 19 | 21 | 25 | 26 | [12 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | [12 empty] | 1
2
| 1 | 1 | 12 | 11 | 13 | < 20 | 11 | 12.1693 | NA | 13 | 14 | 11 | [12 empty] | 1
| 1 | 2 | 12.2 | 11 | 12 | < 20 | 11 | 12.5886 | NA | 12 | 14 | 12 | [12 empty] | 1
| 1 | 3 | 11.9 | 11 | 13 | < 20 | 12 | 11.9569 | NA | 12 | 12 | 11 | [12 empty] | 16
| 1 | 4 | 12.3 | 10 | 12 | < 20 | 11 | 11.865 | NA | 12 | 10 | 11 | [12 empty] | 11.793
| 1 | 5 | 12.2 | 10 | 13 | < 20 | 11 | 12.123 | NA | 11 | 13 | 11 | [12 empty] | 70
| 1 | 6 | 11.8 | 10 | 12 | < 20 | 11 | 11.9602 | NA | 13 | 14 | 11
| Mean |  | 12.0667 | 10.5 | 12.5 |  | 11.1667 | 12.1105 |  | 12.1667 | 12.8333 | 11.1667
| Median |  | 12.1 | 10.5 | 12.5 |  | 11 | 12.0416 |  | 12 | 13.5 | 11
| Std Dev. |  | 0.19664 | 0.54772 | 0.54772 |  | 0.40825 | 0.26023 |  | 0.75277 | 1.6021 | 0.40825
| Rel.Std.Dev. |  | 1.63% | 5.216% | 4.382% |  | 3.656% | 2.149% |  | 6.187% | 12.48% | 3.656%
| PDM3 |  | 2.321% | -10.96% | 5.995% |  | -5.311% | 2.693% |  | 3.169% | 8.822% | -5.311%
| Z-Score (Absolute) |  | 0 | 1.38 | 0.38 | 1.82 | 0.79 | 0.04 | NA | 0.09 | 0.67 | 0.79
| Analytical results for Se in OREAS 555b (Indicative Value 10.5 ppm) | [23 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 25 | 26 | [16 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [16 empty] | 1
1
| 1 | 1 | < 20 | 17 | 7 | < 20 | 11 | < 20 | [16 empty] | 1
| 1 | 2 | < 20 | 14 | 8 | < 20 | 8 | < 20 | [16 empty] | 19
| 1 | 3 | < 20 | 14 | 8 | < 20 | 7 | < 20 | [16 empty] | 16
| 1 | 4 | < 20 | 16 | 7 | < 20 | 11 | < 20 | [16 empty] | 10.5
| 1 | 5 | < 20 | 14 | 7 | < 20 | 12 | < 20 | [16 empty] | 25
| 1 | 6 | < 20 | 23 | 6 | < 20 | 7 | < 20
| Mean |  |  | 16.3333 | 7.1667 |  | 9.3333
| Median |  |  | 15 | 7 |  | 9.5
| Std Dev. |  |  | 3.5024 | 0.75277 |  | 2.2509
| Rel.Std.Dev. |  |  | 21.44% | 10.5% |  | 24.12%
| PDM3 |  |  | 55.56% | -31.75% |  | -11.11%
| Z-Score (Absolute) |  | 0 | 12.81 | 5.73 | 0 | 1.35 | 0
| Analytical results for Si in OREAS 555b (Certified Value 24.62 wt.%) | [23 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 18 | 19 | 21 | 25 | 26 | 28 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 2
3
| 1 | 1 | 25.99 | 24.34 | 25.1 | 25 | 23.8 | 24.21 | 24.82 | 23.37 | 25.01 | 24.68 | 25.5293 | 35.0477 | 26.8192 | 23.7285 | 24.93 | 19.18 | 24.8 | 24.5 |  |  |  |  | 1
| 1 | 2 | 24.2 | 24.18 | 25 | 25.4 | 23.8 | 24.21 | 25.2 | 22.63 | 24.68 | 24.68 | 25.631 | 35.9705 | 24.7568 | 25.1262 | 24.72 | 12.06 | 25.4 | 24.6 |  |  |  |  | #N/A
| 1 | 3 | 27.14 | 24.16 | 24.8 | 24 | 23.9 | 24.12 | 24.59 | 22.77 | 24.59 | 24.59 | 25.7175 | 35.8359 | 24.0476 | 24.3578 | 24.78 | 5.79 | 24.7 | 24 |  |  |  |  | 16
| 1 | 4 | 26.28 | 24.14 | 24.9 | 24.6 | 24 | 24.21 | 24.82 | 24.12 | 24.12 | 24.5 | 25.5962 | 35.0031 | 23.9487 | 24.8243 | 24.55 | 3.5 | 25.1 | 24.2 |  |  |  |  | 24.6187
| 1 | 5 | 25.88 | 24.18 | 23.8 | 24.8 | 24 | 24.21 | 24.68 | 21.83 | 24.31 | 24.68 | 25.7485 | 35.7073 | 24.4095 | 24.6972 | 24.62 | 12.79 | 25 | 24.2 |  |  |  |  | 71
| 1 | 6 | 22.33 | 24.18 | 24.6 | 25.1 | 23.8 | 24.21 | 24.92 | 23.33 | 24.68 | 24.68 | 25.5796 | 35.1133 | 24.6889 | 24.9085 | 24.82 | 1.91 | 24.7 | 24.6
| Mean |  | 25.3033 | 24.1967 | 24.7 | 24.8167 | 23.8833 | 24.195 | 24.8383 | 23.0083 | 24.565 | 24.635 | 25.6337 | 35.4463 | 24.7785 | 24.6071 | 24.7367 | 9.205 | 24.95 | 24.35
| Median |  | 25.935 | 24.18 | 24.85 | 24.9 | 23.85 | 24.21 | 24.82 | 23.05 | 24.635 | 24.68 | 25.6136 | 35.4103 | 24.5492 | 24.7608 | 24.75 | 8.925 | 24.9 | 24.35
| Std Dev. |  | 1.7429 | 0.072019 | 0.47329 | 0.48339 | 0.098319 | 0.036742 | 0.21208 | 0.78211 | 0.31246 | 0.075299 | 0.084182 | 0.43838 | 1.0518 | 0.49994 | 0.13779 | 6.602 | 0.27386 | 0.251
| Rel.Std.Dev. |  | 6.888% | 0.2976% | 1.916% | 1.948% | 0.4117% | 0.1519% | 0.8538% | 3.399% | 1.272% | 0.3057% | 0.3284% | 1.237% | 4.245% | 2.032% | 0.557% | 71.72% | 1.098% | 1.031%
| PDM3 |  | 2.781% | -1.714% | 0.3302% | 0.804% | -2.987% | -1.721% | 0.8921% | -6.541% | -0.2182% | 0.06613% | 4.123% | 43.98% | 0.6489% | -0.04727% | 0.4791% | -62.61% | 1.346% | -1.092%
| Z-Score (Absolute) |  | 1.43 | 1.06 | 0.07 | 0.33 | 1.76 | 1.05 | 0.38 | 3.73 | 0.23 | 0.07 | 2.17 | 24.23 | 0.25 | 0.14 | 0.16 | 34.76 | 0.63 | 0.71
| Analytical results for Sm in OREAS 555b (Certified Value 9.07 ppm) | [23 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 8.9 | 9 | 9.5354 | 9 | 8.3 | 12.3 | 9.38 | [15 empty] | 1
| 1 | 2 | 8.3 | 9.5 | 9.5075 | 8 | 8.3 | 10.7 | 9.13 | [15 empty] | 3
| 1 | 3 | 7.9 | 9.5 | 9.3266 | 7 | 8.4 | 10.9 | 9.16 | [15 empty] | 16
| 1 | 4 | 7.8 | 8.9 | 9.2854 | 10 | 7.6 | 8.6 | 9.51 | [15 empty] | 9.0676
| 1 | 5 | 8.4 | 8.8 | 8.5407 | 7 | 8 | 11.6 | 9.06 | [15 empty] | 72
| 1 | 6 | 8.9 | 9.2 | 9.3778 | 8 | 8.2 | 9.6 | 9.04
| Mean |  | 8.3667 | 9.15 | 9.2622 | 8.1667 | 8.1333 | 10.6167 | 9.2133
| Median |  | 8.35 | 9.1 | 9.3522 | 8 | 8.25 | 10.8 | 9.145
| Std Dev. |  | 0.47188 | 0.30166 | 0.36698 | 1.169 | 0.29439 | 1.3408 | 0.18928
| Rel.Std.Dev. |  | 5.64% | 3.297% | 3.962% | 14.31% | 3.62% | 12.63% | 2.054%
| PDM3 |  | -7.73% | 0.9088% | 2.146% | -9.936% | -10.3% | 17.08% | 1.607%
| Z-Score (Absolute) |  | 1.23 | 0.05 | 0.12 | NA | 1.58 | 2.16 | 0.05
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 555b (Indicative Value 3.47 ppm) | [23 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 15 | 18 | 19 | 21 | 25 | 26 | 27 | 28 | [9 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | [9 empty] | 2
2
| 1 | 1 | < 5 | 14 | < 100 | 4.2 | 4 | < 100 | 2.7449 | < 5 | < 50 | 3 | < 100 | < 50 | < 100 | [9 empty] | 1
| 1 | 2 | < 5 | 20 | < 100 | 4.6 | 4 | < 100 | 3.0886 | < 5 | < 50 | 3 | < 100 | < 50 | < 100 | [9 empty] | 20
| 1 | 3 | < 5 | 17 | < 100 | 3.7 | 4 | < 100 | 2.8171 | < 5 | < 50 | 3 | < 100 | 80 | < 100 | [9 empty] | 16
| 1 | 4 | < 5 | 18 | < 100 | 4.1 | 4 | < 100 | 2.9329 | < 5 | < 50 | 2 | < 100 | < 50 | < 100 | [9 empty] | 3.4735
| 1 | 5 | < 5 | 16 | < 100 | 3.6 | 4 | < 100 | 3.0107 | < 5 | < 50 | 3 | < 100 | < 50 | < 100 | [9 empty] | 26
| 1 | 6 | < 5 | 14 | < 100 | 4.4 | 4 | < 100 | 3.1688 | < 5 | < 50 | 3 | < 100 | < 50 | < 100
| Mean |  |  | 16.5 |  | 4.1 | 4 |  | 2.9605 |  |  | 2.8333 |  | 80
| Median |  |  | 16.5 |  | 4.15 | 4 |  | 2.9718 |  |  | 3 |  | 80
| Std Dev. |  |  | 2.3452 |  | 0.38987 | 0 |  | 0.16133 |  |  | 0.40825
| Rel.Std.Dev. |  |  | 14.21% |  | 9.509% | 0% |  | 5.449% |  |  | 14.41%
| PDM3 |  |  | 375% |  | 18.04% | 15.16% |  | -14.77% |  |  | -18.43% |  | 2203%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.61 | 0.6 | 0.6 | 1.61 | 0.65 | 0.67 | 0.41 | 0.66 | 1.61 | 0.85 | 1.61
| Analytical results for Sr in OREAS 555b (Certified Value 75 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | [6 empty] | 0
1
| 1 | 1 | 64.8 | 80 | 68.4 | 78 | 86 | 72 | 75.3875 | 102 | 74.7 | 71.5214 | NA | 80 | 70 | 35 | 75 | 75 | [6 empty] | 1
| 1 | 2 | 65.8 | 70 | 67.1 | 79 | 86 | 71 | 79.7774 | 103 | 73.7 | 76.9342 | NA | 80 | 75 | < 10 | 74 | 77 | [6 empty] | 32
| 1 | 3 | 65.1 | 80 | 67.9 | 77 | 82 | 72 | 75.5418 | 105 | 75.3 | 77.9601 | NA | 80 | 70 | < 10 | 76 | 78 | [6 empty] | 16
| 1 | 4 | 65 | 80 | 67.7 | 80 | 85 | 73 | 78.3184 | 107 | 73.7 | 70.3956 | NA | 80 | 65 | 61 | 73 | 78 | [6 empty] | 74.9489
| 1 | 5 | 64.5 | 80 | 68.6 | 76 | 83 | 72 | 77.7924 | 104 | 74.9 | 69.1681 | NA | 80 | 75 | 19 | 74 | 76 | [6 empty] | 73
| 1 | 6 | 64.3 | 80 | 67.8 | 80 | 86 | 74 | 78.6045 | 108 | 74.3 | 75.0124 | NA | 80 | 75 | 52 | 72 | 76
| Mean |  | 64.9167 | 78.3333 | 67.9167 | 78.3333 | 84.6667 | 72.3333 | 77.5703 | 104.8333 | 74.4333 | 73.4986 |  | 80 | 71.6667 | 41.75 | 74 | 76.6667
| Median |  | 64.9 | 80 | 67.85 | 78.5 | 85.5 | 72 | 78.0554 | 104.5 | 74.5 | 73.2669 |  | 80 | 72.5 | 43.5 | 74 | 76.5
| Std Dev. |  | 0.52694 | 4.0825 | 0.53448 | 1.633 | 1.7512 | 1.0328 | 1.7567 | 2.3166 | 0.6532 | 3.6412 |  | 0 | 4.0825 | 18.6078 | 1.4142 | 1.2111
| Rel.Std.Dev. |  | 0.8117% | 5.212% | 0.787% | 2.085% | 2.068% | 1.428% | 2.265% | 2.21% | 0.8776% | 4.954% |  | 0% | 5.696% | 44.57% | 1.911% | 1.58%
| PDM3 |  | -13.39% | 4.516% | -9.383% | 4.516% | 12.97% | -3.49% | 3.498% | 39.87% | -0.6879% | -1.935% |  | 6.739% | -4.379% | -44.3% | -1.266% | 2.292%
| Z-Score (Absolute) |  | 1.65 | 0.67 | 1.13 | 0.67 | 1.77 | 0.36 | 0.54 | 5.26 | 0 | 0.16 | NA | 0.96 | 0.48 | 7.77 | 0.07 | 0.39
| Analytical results for Ta in OREAS 555b (Certified Value 1.96 ppm) | [23 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 15 | 18 | 19 | 25 | 26 | 28 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | [12 empty] | 2
3
| 1 | 1 | 2 | 1.9 | 2.3 | 1.9 | 1.8148 | 2.0211 | NA | 3 | 2 | < 5 | [12 empty] | 1
| 1 | 2 | 2 | 2 | 2.6 | 1.9 | 2.7524 | 1.8929 | NA | 2 | 1.8 | < 5 | [12 empty] | 5
| 1 | 3 | 3 | 2 | 2.2 | 1.9 | 2.4265 | 1.8363 | NA | 2 | 1.9 | < 5 | [12 empty] | 16
| 1 | 4 | 2 | 1.9 | 2.2 | 1.9 | 2.399 | 1.8133 | NA | 2 | 1.9 | < 5 | [12 empty] | 1.9644
| 1 | 5 | 2 | 1.9 | 2.1 | 1.8 | 1.8918 | 1.8268 | NA | 2 | 1.9 | < 5 | [12 empty] | 74
| 1 | 6 | < 2 | 2.1 | 3.2 | 1.9 | 2.705 | 1.9407 | NA | 2 | 1.8 | < 5
| Mean |  | 2.2 | 1.9667 | 2.4333 | 1.8833 | 2.3316 | 1.8885 |  | 2.1667 | 1.8833
| Median |  | 2 | 1.95 | 2.25 | 1.9 | 2.4127 | 1.8646 |  | 2 | 1.9
| Std Dev. |  | 0.44721 | 0.08165 | 0.41312 | 0.040825 | 0.39762 | 0.080645 |  | 0.40825 | 0.075277
| Rel.Std.Dev. |  | 20.33% | 4.152% | 16.98% | 2.168% | 17.05% | 4.27% |  | 18.84% | 3.997%
| PDM3 |  | 12% | 0.1168% | 23.87% | -4.125% | 18.69% | -3.861% |  | 10.3% | -4.125%
| Z-Score (Absolute) |  | NA | 0 | 3.78 | 0.67 | 2.95 | 0.63 | NA | NA | 0.67 | 4.32
Results from laboratories 2 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 555b (Certified Value 1.19 ppm) | [23 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 1.2 | 1.12 | 1.1861 | NA | 1.05 | 1.49 | 1.26 | [15 empty] | 1
| 1 | 2 | 1.3 | 1.12 | 1.2063 | NA | 1.2 | 1.22 | 1.23 | [15 empty] | 6
| 1 | 3 | 1.2 | 1.14 | 1.0772 | NA | 1.14 | 1.34 | 1.27 | [15 empty] | 16
| 1 | 4 | 1.2 | 1.13 | 1.1068 | NA | 1.04 | 1.02 | 1.31 | [15 empty] | 1.1933
| 1 | 5 | 1.2 | 1.11 | 1.1131 | NA | 1.12 | 1.41 | 1.29 | [15 empty] | 75
| 1 | 6 | 1.2 | 1.14 | 1.2192 | NA | 1.11 | 1.21 | 1.28
| Mean |  | 1.2167 | 1.1267 | 1.1514 |  | 1.11 | 1.2817 | 1.2733
| Median |  | 1.2 | 1.125 | 1.1496 |  | 1.115 | 1.28 | 1.275
| Std Dev. |  | 0.040825 | 0.012111 | 0.059632 |  | 0.05933 | 0.16774 | 0.027325
| Rel.Std.Dev. |  | 3.355% | 1.075% | 5.179% |  | 5.345% | 13.09% | 2.146%
| PDM3 |  | 1.959% | -5.584% | -3.508% |  | -6.98% | 7.406% | 6.707%
| Z-Score (Absolute) |  | 0.33 | 0.59 | 0.33 | NA | 0.76 | 1 | 0.92
| Analytical results for Te in OREAS 555b (Indicative Value < 2 ppm) | [23 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 25 | 26 | [16 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [16 empty] | 2
2
| 1 | 1 | < 2 | < 6 | < 5 | < 1 | 64 | < 2 | [16 empty] | 1
| 1 | 2 | < 2 | < 6 | < 5 | < 1 | 74 | < 2 | [16 empty] | 7
| 1 | 3 | < 2 | < 6 | < 5 | < 1 | 46 | < 2 | [16 empty] | 16
| 1 | 4 | < 2 | < 6 | < 5 | < 1 | 58 | < 2 | [16 empty] | < 2
| 1 | 5 | < 2 | < 6 | < 5 | < 1 | 78 | < 2 | [16 empty] | 23
| 1 | 6 | < 2 | < 6 | < 5 | < 1 | 54 | < 2
| Mean | [5 empty] | 62.3333
| Median | [5 empty] | 61
| Std Dev. | [5 empty] | 12.16
| Rel.Std.Dev. | [5 empty] | 19.51%
PDM3
| Z-Score (Absolute) |  | 0.51 | 0.84 | 0.51 | 0.84 | 40.85 | 0.51
| Analytical results for Th in OREAS 555b (Certified Value 15.7 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [10 empty] | 1
2
| 1 | 1 | 15.9 | 16 | 15.5 | 15.9 | 15.9446 | 17.3 | 15.8795 | NA | 10.1 | 20 | 15.6 | 14 | [10 empty] | 1
| 1 | 2 | 16.1 | 15.9 | 15.3 | 16.2 | 15.8116 | 18.1 | 15.5056 | NA | 12.2 | 18.8 | 15.7 | 15 | [10 empty] | 8
| 1 | 3 | 16.1 | 15.5 | 14.9 | 16.3 | 15.4458 | 17.9 | 15.7012 | NA | 11.2 | 18.2 | 15.8 | 14 | [10 empty] | 16
| 1 | 4 | 16.1 | 15.5 | 14.9 | 16.1 | 15.6558 | 17.8 | 15.0163 | NA | 12.3 | 15.1 | 15.2 | 15 | [10 empty] | 15.7059
| 1 | 5 | 15.8 | 15.6 | 15.1 | 15.5 | 15.6008 | 17 | 14.9923 | NA | 10.5 | 19.4 | 15.9 | 16 | [10 empty] | 76
| 1 | 6 | 16 | 16.1 | 15.4 | 16.2 | 15.938 | 18.1 | 16.4145 | NA | 11.3 | 17.8 | 15 | 15
| Mean |  | 16 | 15.7667 | 15.1833 | 16.0333 | 15.7328 | 17.7 | 15.5849 |  | 11.2667 | 18.2167 | 15.5333 | 14.8333
| Median |  | 16.05 | 15.75 | 15.2 | 16.15 | 15.7337 | 17.85 | 15.6034 |  | 11.25 | 18.5 | 15.65 | 15
| Std Dev. |  | 0.12649 | 0.26583 | 0.25626 | 0.29439 | 0.19944 | 0.45166 | 0.5421 |  | 0.88242 | 1.7209 | 0.3559 | 0.75277
| Rel.Std.Dev. |  | 0.7906% | 1.686% | 1.688% | 1.836% | 1.268% | 2.552% | 3.478% |  | 7.832% | 9.447% | 2.291% | 5.075%
| PDM3 |  | 1.873% | 0.3872% | -3.327% | 2.085% | 0.1714% | 12.7% | -0.7701% |  | -28.26% | 15.99% | -1.098% | -5.555%
| Z-Score (Absolute) |  | 0.63 | 0.04 | 1.43 | 0.72 | 0.04 | 4.93 | 0.42 | NA | 11.32 | 6.23 | 0.55 | NA
Results from laboratory 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 555b (Certified Value 0.464 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.4 | 0.48 | 0.47 | 0.49 | 0.46 | 0.47 | 0.45 | 0.45 | 0.414 | 0.468 | 0.474 | 0.42795 | 0.5097 | 0.539 | 0.48708 | 0.447 | 0.46 | 0.49 | 0.46 | 0.45 | 0.47 |  | 1
| 1 | 2 | 0.41 | 0.47 | 0.47 | 0.49 | 0.47 | 0.47 | 0.48 | 0.462 | 0.402 | 0.468 | 0.468 | 0.4281 | 0.5025 | 0.531 | 0.47043 | 0.4748 | 0.45 | 0.51 | 0.47 | 0.44 | 0.46 |  | 25
| 1 | 3 | 0.38 | 0.48 | 0.46 | 0.48 | 0.46 | 0.47 | 0.48 | 0.45 | 0.408 | 0.462 | 0.45 | 0.46444 | 0.506 | 0.546 | 0.46671 | 0.4572 | 0.45 | 0.49 | 0.46 | 0.44 | 0.46 |  | 16
| 1 | 4 | 0.4 | 0.47 | 0.46 | 0.48 | 0.46 | 0.47 | 0.45 | 0.462 | 0.426 | 0.456 | 0.474 | 0.44963 | 0.5017 | 0.555 | 0.46917 | 0.4702 | 0.45 | 0.48 | 0.47 | 0.44 | 0.46 |  | 0.46429
| 1 | 5 | 0.4 | 0.47 | 0.48 | 0.46 | 0.46 | 0.47 | 0.45 | 0.45 | 0.384 | 0.462 | 0.474 | 0.46696 | 0.506 | 0.559 | 0.47057 | 0.4698 | 0.44 | 0.49 | 0.46 | 0.43 | 0.46 |  | 77
| 1 | 6 | 0.39 | 0.48 | 0.47 | 0.48 | 0.46 | 0.47 | 0.45 | 0.45 | 0.408 | 0.468 | 0.474 | 0.46599 | 0.5012 | 0.556 | 0.46836 | 0.4697 | 0.44 | 0.53 | 0.46 | 0.44 | 0.46
| Mean |  | 0.39667 | 0.475 | 0.46833 | 0.48 | 0.46167 | 0.47 | 0.46 | 0.454 | 0.407 | 0.464 | 0.469 | 0.45051 | 0.50452 | 0.54767 | 0.47205 | 0.46478 | 0.44833 | 0.49833 | 0.46333 | 0.44 | 0.46167
| Median |  | 0.4 | 0.475 | 0.47 | 0.48 | 0.46 | 0.47 | 0.45 | 0.45 | 0.408 | 0.465 | 0.474 | 0.45704 | 0.50425 | 0.5505 | 0.4698 | 0.46975 | 0.45 | 0.49 | 0.46 | 0.44 | 0.46
| Std Dev. |  | 0.010328 | 0.0054772 | 0.0075277 | 0.010954 | 0.0040825 | 0 | 0.015492 | 0.0061968 | 0.0139 | 0.004899 | 0.0096125 | 0.018525 | 0.0032945 | 0.011021 | 0.0074996 | 0.010514 | 0.0075277 | 0.018348 | 0.005164 | 0.0063246 | 0.0040825
| Rel.Std.Dev. |  | 2.604% | 1.153% | 1.607% | 2.282% | 0.8843% | 0% | 3.368% | 1.365% | 3.415% | 1.056% | 2.05% | 4.112% | 0.653% | 2.012% | 1.589% | 2.262% | 1.679% | 3.682% | 1.115% | 1.437% | 0.8843%
| PDM3 |  | -14.57% | 2.307% | 0.8707% | 3.383% | -0.5652% | 1.23% | -0.9242% | -2.216% | -12.34% | -0.06265% | 1.014% | -2.967% | 8.664% | 17.96% | 1.672% | 0.1061% | -3.437% | 7.332% | -0.2062% | -5.232% | -0.5652%
| Z-Score (Absolute) |  | 4.52 | 0.77 | 0.32 | 1.11 | 0.13 | 0.43 | 0.27 | 0.67 | 3.84 | 0 | 0.34 | 0.88 | 2.76 | 5.67 | 0.57 | 0.08 | 1.03 | 2.34 | 0.02 | 1.59 | 0.13
| Analytical results for Tl in OREAS 555b (Certified Value 1.02 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 15 | 17 | 23 | 25 | 26 | [12 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [12 empty] | 2
3
| 1 | 1 | < 5 | < 2 | 1.1 | 0.9 | 0.8 | 0.93998 | 1.06 | 0.8 | 1.1 | 1 | [12 empty] | 1
| 1 | 2 | < 5 | < 2 | 1 | 1 | 0.7 | 0.9938 | 1.11 | 0.8 | 1 | 1 | [12 empty] | 26
| 1 | 3 | < 5 | < 2 | 1.1 | 1 | 0.6 | 1.0005 | 1.05 | 0.8 | 1 | 1.1 | [12 empty] | 16
| 1 | 4 | < 5 | < 2 | 1 | 1.1 | 0.7 | 0.99727 | 1.09 | 0.8 | 0.9 | 1.1 | [12 empty] | 1.0205
| 1 | 5 | < 5 | < 2 | 1 | 0.9 | 0.7 | 0.91127 | 1.08 | 0.9 | 1 | 1 | [12 empty] | 78
| 1 | 6 | < 5 | < 2 | 1 | 1.1 | 0.7 | 0.94365 | 1.06 | 0.9 | 1 | 1.1
| Mean |  |  |  | 1.0333 | 1 | 0.7 | 0.96442 | 1.075 | 0.83333 | 1 | 1.05
| Median |  |  |  | 1 | 1 | 0.7 | 0.96873 | 1.07 | 0.8 | 1 | 1.05
| Std Dev. |  |  |  | 0.05164 | 0.089443 | 0.063246 | 0.037683 | 0.022583 | 0.05164 | 0.063246 | 0.054772
| Rel.Std.Dev. |  |  |  | 4.997% | 8.944% | 9.035% | 3.907% | 2.101% | 6.197% | 6.325% | 5.216%
| PDM3 |  |  |  | 1.262% | -2.005% | -31.4% | -5.492% | 5.345% | -18.34% | -2.005% | 2.895%
| Z-Score (Absolute) |  | 23.64 | 0 | 0.53 | 0 | 4.73 | 0.56 | 1.18 | 2.63 | 0 | 0.79
| Analytical results for Tm in OREAS 555b (Certified Value 0.57 ppm) | [23 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 0.6 | 0.57 | 0.61702 | NA | 0.51 | 0.62 | 0.58 | [15 empty] | 1
| 1 | 2 | 0.5 | 0.56 | 0.58838 | NA | 0.58 | 0.53 | 0.57 | [15 empty] | 10
| 1 | 3 | 0.5 | 0.6 | 0.55635 | NA | 0.52 | 0.6 | 0.58 | [15 empty] | 16
| 1 | 4 | 0.5 | 0.56 | 0.58726 | NA | 0.47 | 0.5 | 0.59 | [15 empty] | 0.56714
| 1 | 5 | 0.5 | 0.55 | 0.5613 | NA | 0.53 | 0.62 | 0.59 | [15 empty] | 79
| 1 | 6 | 0.5 | 0.58 | 0.60387 | NA | 0.55 | 0.55 | 0.59
| Mean |  | 0.51667 | 0.57 | 0.5857 |  | 0.52667 | 0.57 | 0.58333
| Median |  | 0.5 | 0.565 | 0.58782 |  | 0.525 | 0.575 | 0.585
| Std Dev. |  | 0.040825 | 0.017889 | 0.023568 |  | 0.037238 | 0.050596 | 0.008165
| Rel.Std.Dev. |  | 7.902% | 3.138% | 4.024% |  | 7.071% | 8.877% | 1.4%
| PDM3 |  | -8.899% | 0.5045% | 3.272% |  | -7.136% | 0.5045% | 2.855%
| Z-Score (Absolute) |  | NA | 0 | 0.79 | NA | 2.19 | 0 | 0.67
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 555b (Certified Value 9.51 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 23 | 25 | 26 | 28 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [11 empty] | 2
3
| 1 | 1 | 9.2 | 9.6 | 9.4 | 10.0622 | 10.2 | 8.9996 | 9 | 8 | 11.4 | 9.7 | 9 | [11 empty] | 1
| 1 | 2 | 9.5 | 9.3 | 9.4 | 10.0018 | 10.1 | 8.8289 | 10 | 8 | 8.21 | 9.5 | 9 | [11 empty] | 11
| 1 | 3 | 9.4 | 9 | 9.4 | 9.9962 | 10.3 | 8.6218 | 9 | 8 | 7.37 | 9.3 | 10 | [11 empty] | 16
| 1 | 4 | 9.1 | 9.2 | 9.4 | 10.1852 | 10.2 | 8.747 | 10 | 7 | 3.76 | 9.6 | 10 | [11 empty] | 9.5113
| 1 | 5 | 9.2 | 9.2 | 9.2 | 10.1433 | 10.3 | 8.1713 | 9 | 7 | 8.45 | 9.5 | 9 | [11 empty] | 80
| 1 | 6 | 9.1 | 9.5 | 9.6 | 10.1181 | 10.2 | 8.7746 | 9 | 7 | 6.14 | 9.6 | 9
| Mean |  | 9.25 | 9.3 | 9.4 | 10.0845 | 10.2167 | 8.6905 | 9.3333 | 7.5 | 7.555 | 9.5333 | 9.3333
| Median |  | 9.2 | 9.25 | 9.4 | 10.0902 | 10.2 | 8.7608 | 9 | 7.5 | 7.79 | 9.55 | 9
| Std Dev. |  | 0.16432 | 0.21909 | 0.12649 | 0.077279 | 0.075277 | 0.28258 | 0.5164 | 0.54772 | 2.5476 | 0.13663 | 0.5164
| Rel.Std.Dev. |  | 1.776% | 2.356% | 1.346% | 0.7663% | 0.7368% | 3.252% | 5.533% | 7.303% | 33.72% | 1.433% | 5.533%
| PDM3 |  | -2.747% | -2.221% | -1.17% | 6.027% | 7.416% | -8.629% | -1.871% | -21.15% | -20.57% | 0.232% | -1.871%
| Z-Score (Absolute) |  | 0.16 | 0.08 | 0.08 | 1.18 | 1.39 | 1.06 | NA | NA | 2.87 | 0.29 | NA
Results from laboratories 19, 23 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 555b (Certified Value 379 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 15 | 16 | 17 | 19 | 21 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 351.1 | 360 | 366 | 392 | 421 | 386 | 401.743 | 402 | 370 | NA | 340 | 339 | 379 | 368 | 377 | [7 empty] | 1
| 1 | 2 | 351.8 | 350 | 363 | 390 | 409 | 381 | 403.5085 | 406 | 380 | NA | 330 | 403 | 417 | 378 | 384 | [7 empty] | 38
| 1 | 3 | 340.2 | 360 | 365 | 395 | 412 | 388 | 392.0718 | 407 | 370 | NA | 340 | 369 | 424 | 377 | 382 | [7 empty] | 16
| 1 | 4 | 350.6 | 350 | 360 | 393 | 414 | 392 | 400.8072 | 405 | 370 | NA | 330 | 340 | 404 | 367 | 385 | [7 empty] | 378.7095
| 1 | 5 | 351.7 | 350 | 370 | 382 | 408 | 392 | 391.8872 | 402 | 380 | NA | 330 | 378 | 401 | 372 | 370 | [7 empty] | 81
| 1 | 6 | 339.3 | 350 | 367 | 392 | 407 | 395 | 395.8795 | 406 | 370 | NA | 330 | 410 | 444 | 367 | 378
| Mean |  | 347.45 | 353.3333 | 365.1667 | 390.6667 | 411.8333 | 389 | 397.6495 | 404.6667 | 373.3333 |  | 333.3333 | 373.1667 | 411.5 | 371.5 | 379.3333
| Median |  | 350.85 | 350 | 365.5 | 392 | 410.5 | 390 | 398.3433 | 405.5 | 370 |  | 330 | 373.5 | 410.5 | 370 | 380
| Std Dev. |  | 5.9869 | 5.164 | 3.4303 | 4.5461 | 5.1929 | 5.0596 | 5.0696 | 2.1602 | 5.164 |  | 5.164 | 30.1822 | 22.2059 | 5.01 | 5.5737
| Rel.Std.Dev. |  | 1.723% | 1.462% | 0.9394% | 1.164% | 1.261% | 1.301% | 1.275% | 0.5338% | 1.383% |  | 1.549% | 8.088% | 5.396% | 1.349% | 1.469%
| PDM3 |  | -8.254% | -6.701% | -3.576% | 3.157% | 8.747% | 2.717% | 5.001% | 6.854% | -1.42% |  | -11.98% | -1.464% | 8.658% | -1.904% | 0.1647%
| Z-Score (Absolute) |  | 1.09 | 0.87 | 0.42 | 0.54 | 1.34 | 0.48 | 0.81 | 1.07 | 0.11 | NA | 1.63 | 0.12 | 1.33 | 0.18 | 0.11
| Analytical results for W in OREAS 555b (Certified Value 2.94 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 15 | 19 | 21 | 25 | 26 | 28 | [11 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | [11 empty] | 2
2
| 1 | 1 | < 5 | < 10 | 4 | 3.4 | < 5 | 3.1351 | NA | < 50 | < 1 | 3 | < 5 | [11 empty] | 1
| 1 | 2 | < 5 | < 10 | 4 | 3.6 | < 5 | 2.884 | NA | < 50 | < 1 | 3 | < 5 | [11 empty] | 12
| 1 | 3 | < 5 | < 10 | 3 | 5.2 | < 5 | 2.8512 | NA | < 50 | < 1 | 3 | < 5 | [11 empty] | 16
| 1 | 4 | < 5 | < 10 | 3 | 4 | < 5 | 2.9899 | NA | < 50 | 2 | 3 | < 5 | [11 empty] | 2.9367
| 1 | 5 | < 5 | < 10 | 4 | 3.5 | < 5 | 3.2656 | NA | < 50 | 2 | 3 | < 5 | [11 empty] | 82
| 1 | 6 | < 5 | < 10 | 4 | 4 | < 5 | 3.2754 | NA | < 50 | 2 | 3 | < 5
| Mean |  |  |  | 3.6667 | 3.95 |  | 3.0669 |  |  | 2 | 3
| Median |  |  |  | 4 | 3.8 |  | 3.0625 |  |  | 2 | 3
| Std Dev. |  |  |  | 0.5164 | 0.66257 |  | 0.18627 |  |  | 0 | 0
| Rel.Std.Dev. |  |  |  | 14.08% | 16.77% |  | 6.074% |  |  | 0% | 0%
| PDM3 |  |  |  | 24.86% | 34.5% |  | 4.432% |  |  | -31.9% | 2.155%
| Z-Score (Absolute) |  | 0.62 | 2.27 | 0.73 | 1.06 | 0.62 | 0.04 | NA | 25.39 | 2.06 | 0.04 | 0.62
| Analytical results for Y in OREAS 555b (Certified Value 36.5 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 28 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | [9 empty] | 1
2
| 1 | 1 | 36.3 | 36.3 | 39.5 | 35 | 30.5304 | 39 | 35.681 | 51 | 34 | 29.4 | 40.8 | 37.3 | 39 | [9 empty] | 1
| 1 | 2 | 34.7 | 36.3 | 38.3 | 36.2 | 32.8455 | 38.6 | 36.1728 | 43 | 32 | 30.3 | 37.3 | 37.3 | 37 | [9 empty] | 13
| 1 | 3 | 35.4 | 36.3 | 35.9 | 36.5 | 31.284 | 38.6 | 35.4904 | 50 | 32 | 29.6 | 38.7 | 37.6 | 38 | [9 empty] | 16
| 1 | 4 | 34.9 | 37.9 | 36.5 | 34.8 | 32.5242 | 39.7 | 35.5395 | 41 | 34 | 30.6 | 30.9 | 36.5 | 39 | [9 empty] | 36.5396
| 1 | 5 | 37.4 | 36.4 | 37.9 | 35.5 | 31.5797 | 39 | 35.6474 | 37 | 32 | 30.5 | 39.3 | 37 | 38 | [9 empty] | 83
| 1 | 6 | 36.3 | 36.9 | 36.9 | 35.3 | 31.4776 | 38.7 | 37.2637 | 43 | 34 | 30 | 36.2 | 36.7 | 38
| Mean |  | 35.8333 | 36.6833 | 37.5 | 35.55 | 31.7069 | 38.9333 | 35.9658 | 44.1667 | 33 | 30.0667 | 37.2 | 37.0667 | 38.1667
| Median |  | 35.85 | 36.35 | 37.4 | 35.4 | 31.5287 | 38.85 | 35.6642 | 43 | 33 | 30.15 | 38 | 37.15 | 38
| Std Dev. |  | 1.0231 | 0.64005 | 1.3206 | 0.67157 | 0.84807 | 0.41793 | 0.68083 | 5.3821 | 1.0954 | 0.48854 | 3.4733 | 0.41312 | 0.75277
| Rel.Std.Dev. |  | 2.855% | 1.745% | 3.522% | 1.889% | 2.675% | 1.073% | 1.893% | 12.19% | 3.32% | 1.625% | 9.337% | 1.115% | 1.972%
| PDM3 |  | -1.933% | 0.3933% | 2.628% | -2.708% | -13.23% | 6.551% | -1.57% | 20.87% | -9.687% | -17.71% | 1.807% | 1.442% | 4.453%
| Z-Score (Absolute) |  | 0.51 | 0 | 0.49 | 0.67 | 2.96 | 1.34 | 0.43 | 4.45 | 2.19 | 3.94 | 0.31 | 0.23 | 0.88
| Analytical results for Yb in OREAS 555b (Certified Value 3.49 ppm) | [23 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | 18 | 19 | 23 | 25 | 28 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [15 empty] | 2
3
| 1 | 1 | 3.6 | 3.7 | 3.5771 | 3 | 3.3 | 3.4 | 3.68 | [15 empty] | 1
| 1 | 2 | 3.6 | 3.8 | 3.5648 | 4 | 3.6 | 3.3 | 3.56 | [15 empty] | 14
| 1 | 3 | 3.4 | 3.7 | 3.4099 | 4 | 3.6 | 3.1 | 3.67 | [15 empty] | 16
| 1 | 4 | 3.8 | 3.7 | 3.6433 | 4 | 3.1 | 3 | 3.75 | [15 empty] | 3.4947
| 1 | 5 | 3.2 | 3.6 | 3.5118 | 4 | 3.3 | 3.5 | 3.65 | [15 empty] | 84
| 1 | 6 | 3.3 | 3.6 | 3.5424 | 3 | 3.4 | 3.1 | 3.55
| Mean |  | 3.4833 | 3.6833 | 3.5415 | 3.6667 | 3.3833 | 3.2333 | 3.6433
| Median |  | 3.5 | 3.7 | 3.5536 | 4 | 3.35 | 3.2 | 3.66
| Std Dev. |  | 0.22286 | 0.075277 | 0.077948 | 0.5164 | 0.19408 | 0.19664 | 0.076333
| Rel.Std.Dev. |  | 6.398% | 2.044% | 2.201% | 14.08% | 5.736% | 6.082% | 2.095%
| PDM3 |  | -0.3253% | 5.398% | 1.34% | 4.921% | -3.187% | -7.479% | 4.253%
| Z-Score (Absolute) |  | 0.15 | 0.89 | 0.15 | NA | 0.67 | 1.45 | 0.68
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 555b (Certified Value 31.7 ppm) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 1
1
| 1 | 1 | 38.8 | 56 | 27.7 | 41.37 | 50 | 28 | < 20 | 60 | 30 | 20 | 50 | 24.3639 | 30 | 31.2008 | NA | 30 | 30 | 28 | 43 | < 100 | < 100 |  | 1
| 1 | 2 | 28.6 | 56 | 27.7 | 40.15 | 40 | 27 | < 20 | 20 | 30 | 20 | < 20 | 27.6343 | 30 | 29.2015 | NA | 30 | 30 | 19 | 36 | < 100 | < 100 |  | 15
| 1 | 3 | 30.5 | 60 | 27.7 | 33.47 | 40 | 29 | < 20 | 20 | 40 | 20 | < 20 | 27.8407 | 30 | 30.5428 | NA | 30 | 40 | 17 | 33 | < 100 | < 100 |  | 16
| 1 | 4 | 34.2 | 56 | 27 | 36.2 | 40 | 29 | < 20 | 20 | 20 | < 20 | 380 | 23.9628 | 30 | 30.0497 | NA | 40 | 35 | 12 | 36 | < 100 | < 100 |  | 31.6644
| 1 | 5 | 31.2 | 55 | 27.8 | 44.29 | 50 | 29 | < 20 | 30 | 20 | 20 | 40 | 29.0325 | 30 | 31.5422 | NA | 40 | 35 | 17 | 36 | < 100 | < 100 |  | 85
| 1 | 6 | 45.9 | 63 | 27.9 | 33.5 | 50 | 28 | 20 | 20 | 30 | < 20 | 30 | 25.7636 | 30 | 30.9772 | NA | 50 | 25 | 15 | 33 | < 100 | < 100
| Mean |  | 34.8667 | 57.6667 | 27.6333 | 38.1633 | 45 | 28.3333 | 20 | 28.3333 | 28.3333 | 20 | 125 | 26.4329 | 30 | 30.5857 |  | 36.6667 | 32.5 | 18 | 36.1667
| Median |  | 32.7 | 56 | 27.7 | 38.175 | 45 | 28.5 | 20 | 20 | 30 | 20 | 45 | 26.6989 | 30 | 30.76 |  | 35 | 32.5 | 17 | 36
| Std Dev. |  | 6.472 | 3.1411 | 0.32042 | 4.4586 | 5.4772 | 0.8165 |  | 16.0208 | 7.5277 | 0 | 170.196 | 2.0501 | 0 | 0.85457 |  | 8.165 | 5.244 | 5.4406 | 3.656
| Rel.Std.Dev. |  | 18.56% | 5.447% | 1.16% | 11.68% | 12.17% | 2.882% |  | 56.54% | 26.57% | 0% | 136.2% | 7.756% | 0% | 2.794% |  | 22.27% | 16.14% | 30.23% | 10.11%
| PDM3 |  | 10.11% | 82.12% | -12.73% | 20.52% | 42.12% | -10.52% | -36.84% | -10.52% | -10.52% | -36.84% | 294.8% | -16.52% | -5.256% | -3.407% |  | 15.8% | 2.639% | -43.15% | 14.22%
| Z-Score (Absolute) |  | 0.14 | 2.84 | 0.72 | 0.53 | NA | 0.63 | NA | NA | NA | NA | NA | 0.86 | NA | 0.37 | NA | NA | 0.14 | 1.86 | 0.29 | 1.93 | 1.93
| Analytical results for Zr in OREAS 555b (Indicative Value 247 ppm) | [23 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 06 | 18 | 19 | 25 | [17 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | [17 empty] | 0
0
| 1 | 1 | 188.2 | 242 | 275.4071 | 226 | 308 | [17 empty] | 1
| 1 | 2 | 185.6 | 246 | 251.8388 | 234 | 239 | [17 empty] | 18
| 1 | 3 | 181.5 | 299 | 247.6486 | 229 | 260 | [17 empty] | 16
| 1 | 4 | 189.1 | 235 | 245.7964 | 237 | 213 | [17 empty] | 246.5314
| 1 | 5 | 185.3 | 243 | 244.287 | 231 | 309 | [17 empty] | 24
| 1 | 6 | 188.1 | 241 | 251.3752 | 231 | 235
| Mean |  | 186.3 | 251 | 252.7255 | 231.3333 | 260.6667
| Median |  | 186.85 | 242.5 | 249.5119 | 231 | 249.5
| Std Dev. |  | 2.8007 | 23.7908 | 11.507 | 3.8297 | 39.9433
| Rel.Std.Dev. |  | 1.503% | 9.478% | 4.553% | 1.655% | 15.32%
| PDM3 |  | -24.43% | 1.813% | 2.513% | -6.165% | 5.734%
| Z-Score (Absolute) |  | 4.51 | 0 | 0.12 | 1.37 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1038: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 555b (Certified Value 0.098 ppm) | [28 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 1.98 | 0.3 | < 0.5 | 0.08 | < 0.3 | 0.1 | 0.09 | < 0.01 | 0.13 | 0.13 | 0.08 | < 5 | 0.9 | < 0.1 | 0.1 | 0.21232 | < 0.3 | 0.2 | 0.2 | 0.2 | 0.1 | 0.1 | < 0.5 | 0.1 |  |  |  | 1
| 1 | 2 | 1.97 | 0.11 | < 0.5 | 0.08 | 0.5 | 0.11 | 0.08 | < 0.01 | 0.11 | 0.08 | 0.08 | < 5 | < 0.2 | < 0.1 | < 0.1 | 0.20811 | < 0.3 | 0.2 | < 0.2 | 0.2 | 0.1 | 0.09 | < 0.5 | 0.09 |  |  |  | 1
| 1 | 3 | 1.96 | 0.36 | < 0.5 | 0.09 | < 0.3 | 0.1 | 0.1 | < 0.01 | 0.09 | 0.11 | 0.1 | < 5 | 0.9 | < 0.1 | < 0.1 | 0.2093 | < 0.3 | 0.2 | 0.2 | 0.2 | 0.09 | 0.09 | < 0.5 | 0.1 |  |  |  | 16
| 1 | 4 | 1.92 | 0.36 | < 0.5 | 0.08 | 0.3 | 0.11 | 0.09 | < 0.01 | 0.1 | 0.13 | 0.08 | < 5 | 2.4 | < 0.1 | < 0.1 | 0.20842 | < 0.3 | 0.2 | < 0.2 | 0.2 | 0.08 | 0.1 | < 0.5 | 0.12 |  |  |  | 0.097963
| 1 | 5 | 1.96 | 0.31 | < 0.5 | 0.09 | 0.3 | 0.11 | 0.1 | 0.01 | 0.09 | 0.11 | 0.12 | < 5 | 2.9 | < 0.1 | < 0.1 | 0.21006 | < 0.3 | 0.2 | 0.2 | 0.2 | 0.1 | 0.08 | < 0.5 | 0.1 |  |  |  | 87
| 1 | 6 | 1.94 | 0.31 | < 0.5 | 0.09 | 0.3 | 0.1 | 0.08 | 0.03 | 0.13 | 0.12 | 0.09 | < 5 | 4.1 | < 0.1 | < 0.1 | 0.20705 | < 0.3 | 0.2 | < 0.2 | 0.2 | 0.09 | 0.1 | < 0.5 | 0.1
| Mean |  | 1.955 | 0.29167 |  | 0.085 | 0.35 | 0.105 | 0.09 | 0.02 | 0.10833 | 0.11333 | 0.091667 |  | 2.24 |  | 0.1 | 0.20921 |  | 0.2 | 0.2 | 0.2 | 0.093333 | 0.093333 |  | 0.10167
| Median |  | 1.96 | 0.31 |  | 0.085 | 0.3 | 0.105 | 0.09 | 0.02 | 0.105 | 0.115 | 0.085 |  | 2.4 |  | 0.1 | 0.20886 |  | 0.2 | 0.2 | 0.2 | 0.095 | 0.095 |  | 0.1
| Std Dev. |  | 0.021679 | 0.092826 |  | 0.0054772 | 0.1 | 0.0054772 | 0.0089443 | 0.014142 | 0.018348 | 0.018619 | 0.016021 |  | 1.3704 |  |  | 0.0018416 |  | 3.0405e-17 | 3.3993e-17 | 3.0405e-17 | 0.008165 | 0.008165 |  | 0.0098319
| Rel.Std.Dev. |  | 1.109% | 31.83% |  | 6.444% | 28.57% | 5.216% | 9.938% | 70.71% | 16.94% | 16.43% | 17.48% |  | 61.18% |  |  | 0.8803% |  | 1.52e-14% | 1.7e-14% | 1.52e-14% | 8.748% | 8.748% |  | 9.671%
| PDM3 |  | 1896% | 197.7% |  | -13.23% | 257.3% | 7.183% | -8.129% | -79.58% | 10.59% | 15.69% | -6.427% |  | 2187% |  | 2.079% | 113.6% |  | 104.2% | 104.2% | 104.2% | -4.726% | -4.726% |  | 3.781%
| Z-Score (Absolute) |  | 38.35 | 3.84 | 2.97 | 0.45 | NA | 0.03 | 0.35 | 2.01 | 0.03 | 0.14 | 0.31 | 49.66 | NA | 1.18 | NA | 2.13 | 0.9 | NA | NA | NA | 0.28 | 0.28 | 2.97 | 0.1
Results from laboratories 5, 14, 17, 20, 21 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 555b (Certified Value 6.4 wt.%) | [28 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.06 | 6.81 | 6.31 | 6.4924 | 6.6 | 6.02 | 6.29 | 6.2 | 6.24 | 6.32 | 6.66 | 6.3706 | 5.663 | 6.4101 | 6.4086 | 6.52 | > 5 | 6.0754 | 6.48 | 6.46 | 7 | 6.78 | 5.99 | 6.3798 | 6.43 | 6.4 |  | 1
| 1 | 2 | 5.99 | 6.23 | 6.34 | 6.4187 | 6.62 | 5.85 | 6.21 | 6.31 | 6.45 | 6.45 | 6.56 | 6.58 | 5.61 | 6.4676 | 6.4773 | 6.58 | > 5 | 6.1269 | 6.44 | 6.51 | 7.04 | 6.73 | 5.87 | 6.2223 | 6.43 | 6.36 |  | #N/A
| 1 | 3 | 5.93 | 6.74 | 6.33 | 6.4323 | 6.47 | 6.02 | 6.38 | 6.38 | 6.45 | 6.5 | 6.63 | 6.5134 | 5.716 | 6.5042 | 6.4719 | 6.65 | > 5 | 6.208 | 6.36 | 6.37 | 7.37 | 6.93 | 6.38 | 6.3047 | 6.4 | 6.2 |  | 16
| 1 | 4 | 6.07 | 7.05 | 6.33 | 6.3172 | 6.47 | 5.81 | 6.27 | 6.34 | 6.49 | 6.42 | 6.58 | 6.5227 | 5.716 | 6.4129 | 6.44 | 6.69 | > 5 | 6.0715 | 6.4 | 6.4 | 7.15 | 6.69 | 6.66 | 6.2609 | 6.4 | 6.42 |  | 6.4043
| 1 | 5 | 5.84 | 7.07 | 6.31 | 6.4182 | 6.42 | 5.94 | 6.24 | 6.4 | 6.43 | 6.39 | 6.63 | 6.5347 | 5.663 | 6.5338 | 6.4989 | 6.77 | > 5 | 6.0784 | 6.26 | 6.35 | 7.38 | 6.88 | 6.49 | 6.336 | 6.45 | 6.31 |  | 88
| 1 | 6 | 5.88 | 6.78 | 6.35 | 6.3856 | 6.43 | 5.89 | 6.16 | 6.25 | 6.27 | 6.55 | 6.61 | 6.5443 | 5.663 | 6.3686 | 6.4123 | 6.66 | > 5 | 6.1046 | 6.49 | 6.39 | 7.46 | 6.69 | 7.04 | 6.2606 | 6.45 | 6.37
| Mean |  | 5.9617 | 6.78 | 6.3283 | 6.4107 | 6.5017 | 5.9217 | 6.2583 | 6.3133 | 6.3883 | 6.4383 | 6.6117 | 6.511 | 5.6718 | 6.4495 | 6.4515 | 6.645 |  | 6.1108 | 6.405 | 6.4133 | 7.2333 | 6.7833 | 6.405 | 6.294 | 6.4267 | 6.3433
| Median |  | 5.96 | 6.795 | 6.33 | 6.4184 | 6.47 | 5.915 | 6.255 | 6.325 | 6.44 | 6.435 | 6.62 | 6.5287 | 5.663 | 6.4402 | 6.4559 | 6.655 |  | 6.0915 | 6.42 | 6.395 | 7.26 | 6.755 | 6.435 | 6.2828 | 6.43 | 6.365
| Std Dev. |  | 0.094534 | 0.30397 | 0.016021 | 0.057678 | 0.086583 | 0.087502 | 0.075211 | 0.076855 | 0.10553 | 0.081343 | 0.03656 | 0.072507 | 0.039897 | 0.063017 | 0.036981 | 0.086891 |  | 0.05214 | 0.086197 | 0.060222 | 0.19511 | 0.10113 | 0.43233 | 0.057606 | 0.022509 | 0.079666
| Rel.Std.Dev. |  | 1.586% | 4.483% | 0.2532% | 0.8997% | 1.332% | 1.478% | 1.202% | 1.217% | 1.652% | 1.263% | 0.553% | 1.114% | 0.7034% | 0.9771% | 0.5732% | 1.308% |  | 0.8532% | 1.346% | 0.939% | 2.697% | 1.491% | 6.75% | 0.9152% | 0.3502% | 1.256%
| PDM3 |  | -6.911% | 5.867% | -1.186% | 0.1011% | 1.521% | -7.535% | -2.279% | -1.42% | -0.2487% | 0.5321% | 3.239% | 1.666% | -11.44% | 0.707% | 0.7377% | 3.759% |  | -4.582% | 0.01159% | 0.1417% | 12.95% | 5.919% | 0.01159% | -1.721% | 0.3499% | -0.9513%
| Z-Score (Absolute) |  | 3.11 | 2.56 | 0.57 | 0 | 0.63 | 3.39 | 1.06 | 0.67 | 0.16 | 0.19 | 1.39 | 0.69 | 5.12 | 0.27 | 0.28 | 1.62 | NA | 2.08 | 0.04 | 0.02 | 5.69 | 2.58 | 0.04 | 0.81 | 0.11 | 0.47
| Analytical results for As in OREAS 555b (Certified Value 64 ppm) | [28 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 41.7 | 61 | 65.8 | 61.9 | 66 | 61.4 | 62.8 | 62.9 | 64.3 | 65.7 | 66.8 | 57.17 | 197 | 64.0763 | 72.84 | 65 | 66.4517 | 63.7 | 67.3 | 65 | 39.7 | 46 | 54.7 | 62 | 34 | 63.4 |  | 1
| 1 | 2 | 35.3 | 65.5 | 65.8 | 59.8 | 67 | 62.9 | 65.4 | 64.2 | 68.7 | 64 | 62.7 | 56.57 | 253 | 63.8672 | 74.96 | 65 | 67.045 | 58.2 | 64.2 | 61 | 36 | 44 | 63 | 63 | 33 | 62 |  | 2
| 1 | 3 | 41.76 | 62.9 | 65.4 | 63.3 | 67.7 | 62.5 | 65.7 | 65.1 | 66.9 | 68.2 | 65.2 | 56.84 | 159 | 63.5003 | 73.49 | 65 | 66.799 | 61 | 67.2 | 64 | 35.6 | 42 | 56.7 | 64 | 35 | 63.8 |  | 16
| 1 | 4 | 33.72 | 63 | 67.1 | 63.3 | 50.8 | 62.3 | 63.1 | 60.7 | 68.7 | 65.9 | 61.1 | 56.96 | 137 | 63.8638 | 69.92 | 66 | 66.7404 | 60.9 | 65.9 | 61 | 35.2 | 43 | 57.8 | 61 | 33 | 64.6 |  | 63.9806
| 1 | 5 | 37.44 | 63.1 | 66 | 63.7 | 65.4 | 65.7 | 64.2 | 65.7 | 64.3 | 63.5 | 64.3 | 56.97 | 131 | 63.4338 | 70.23 | 64 | 67.272 | 57.5 | 67.2 | 66 | 35.5 | 44 | 59 | 64 | 34 | 62.4 |  | 89
| 1 | 6 | 37.79 | 59 | 66.5 | 64.2 | 66.2 | 66.8 | 62.7 | 64.5 | 68.4 | 67.1 | 63.5 | 56.88 | 150 | 63.5686 | 70.83 | 64 | 65.336 | 59.6 | 66.2 | 63 | 36 | 45 | 56.5 | 62 | 36 | 64.2
| Mean |  | 37.9517 | 62.4167 | 66.1 | 62.7 | 63.85 | 63.6 | 63.9833 | 63.85 | 66.8833 | 65.7333 | 63.9333 | 56.8983 | 171.1667 | 63.7184 | 72.045 | 64.8333 | 66.6073 | 60.15 | 66.3333 | 63.3333 | 36.3333 | 44 | 57.95 | 62.6667 | 34.1667 | 63.4
| Median |  | 37.615 | 62.95 | 65.9 | 63.3 | 66.1 | 62.7 | 63.65 | 64.35 | 67.65 | 65.8 | 63.9 | 56.92 | 154.5 | 63.7162 | 71.835 | 65 | 66.7697 | 60.25 | 66.7 | 63.5 | 35.8 | 44 | 57.25 | 62.5 | 34 | 63.6
| Std Dev. |  | 3.2796 | 2.2013 | 0.60663 | 1.6137 | 6.4435 | 2.1392 | 1.3288 | 1.8086 | 2.1094 | 1.7874 | 1.9846 | 0.19712 | 46.3483 | 0.25391 | 2.025 | 0.75277 | 0.68239 | 2.2367 | 1.1994 | 2.0656 | 1.6777 | 1.4142 | 2.8585 | 1.2111 | 1.169 | 1.0198
| Rel.Std.Dev. |  | 8.642% | 3.527% | 0.9177% | 2.574% | 10.09% | 3.363% | 2.077% | 2.833% | 3.154% | 2.719% | 3.104% | 0.3464% | 27.08% | 0.3985% | 2.811% | 1.161% | 1.024% | 3.719% | 1.808% | 3.261% | 4.618% | 3.214% | 4.933% | 1.933% | 3.422% | 1.609%
| PDM3 |  | -40.68% | -2.444% | 3.313% | -2.002% | -0.2041% | -0.5949% | 0.004265% | -0.2041% | 4.537% | 2.739% | -0.07388% | -11.07% | 167.5% | -0.4099% | 12.6% | 1.333% | 4.106% | -5.987% | 3.677% | -1.012% | -43.21% | -31.23% | -9.426% | -2.054% | -46.6% | -0.9075%
| Z-Score (Absolute) |  | 7.68 | 0.37 | 0.73 | 0.29 | 0.06 | 0.02 | 0.1 | 0.06 | 0.96 | 0.62 | 0.08 | 2.02 | 32.11 | 0.02 | 2.5 | 0.35 | 0.88 | 1.05 | 0.8 | 0.1 | 8.16 | 5.87 | 1.71 | 0.3 | 8.81 | 0.08
| Analytical results for Ba in OREAS 555b (Certified Value 254 ppm) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 255.3 | 45 | 245 | 265.9 | 261 | 268 | 212.5455 | 250 | 250 | 260 | 250 | 270 | 230 | 246 | 268.1382 | 253.22 | 247 | 228.6822 | 251 | 256 | 468 | 269 | 201 | 253 | 254 | 250 |  | 1
| 1 | 2 | 253 | 73 | 244 | 260.3 | 262 | 268 | 214.3341 | 250 | 260 | 270 | 250 | 270 | 227 | 244 | 263.4534 | 255.46 | 247 | 226.9583 | 251 | 264 | 488 | 266 | 198 | 257 | 243 | 247 |  | 30
| 1 | 3 | 252.6 | 43 | 240 | 269.2 | 264 | 271 | 217.9702 | 250 | 260 | 270 | 260 | 270 | 229 | 241 | 261.0658 | 259.51 | 254 | 216.8812 | 254 | 254 | 478 | 263 | 205 | 257 | 260 | 239 |  | 16
| 1 | 4 | 258.1 | 54 | 245 | 269.5 | 263 | 269 | 206.4579 | 250 | 250 | 270 | 250 | 270 | 228 | 241 | 269.093 | 260.41 | 253 | 227.8005 | 250 | 263 | 481 | 268 | 214 | 252 | 255 | 243 |  | 254.294
| 1 | 5 | 247.8 | 37 | 242 | 271.5 | 257 | 275 | 204.2763 | 250 | 250 | 270 | 240 | 270 | 230 | 243 | 262.5036 | 257.45 | 259 | 229.2508 | 249 | 262 | 510 | 280 | 205 | 260 | 260 | 246 |  | 90
| 1 | 6 | 250.2 | 41 | 242 | 277.7 | 256 | 273 | 209.5641 | 250 | 250 | 260 | 260 | 270 | 231 | 242 | 265.9258 | 252.26 | 255 | 218.0481 | 250 | 254 | 473 | 275 | 211 | 255 | 259 | 234
| Mean |  | 252.8333 | 48.8333 | 243 | 269.0167 | 260.5 | 270.6667 | 210.858 | 250 | 253.3333 | 266.6667 | 251.6667 | 270 | 229.1667 | 242.8333 | 265.03 | 256.385 | 252.5 | 224.6035 | 250.8333 | 258.8333 | 483 | 270.1667 | 205.6667 | 255.6667 | 255.1667 | 243.1667
| Median |  | 252.8 | 44 | 243 | 269.35 | 261.5 | 270 | 211.0548 | 250 | 250 | 270 | 250 | 270 | 229.5 | 242.5 | 264.6896 | 256.455 | 253.5 | 227.3794 | 250.5 | 259 | 479.5 | 268.5 | 205 | 256 | 257 | 244.5
| Std Dev. |  | 3.6379 | 13.1212 | 2 | 5.7884 | 3.2711 | 2.8752 | 5.0964 | 0 | 5.164 | 5.164 | 7.5277 | 0 | 1.472 | 1.9408 | 3.2111 | 3.3163 | 4.7223 | 5.5966 | 1.7224 | 4.6655 | 14.8862 | 6.2423 | 5.9889 | 2.9439 | 6.4936 | 5.8452
| Rel.Std.Dev. |  | 1.439% | 26.87% | 0.823% | 2.152% | 1.256% | 1.062% | 2.417% | 0% | 2.038% | 1.936% | 2.991% | 0% | 0.6423% | 0.7992% | 1.212% | 1.293% | 1.87% | 2.492% | 0.6867% | 1.803% | 3.082% | 2.311% | 2.912% | 1.151% | 2.545% | 2.404%
| PDM3 |  | -0.5744% | -80.8% | -4.441% | 5.79% | 2.44% | 6.438% | -17.08% | -1.689% | -0.3778% | 4.865% | -1.033% | 6.176% | -9.881% | -4.507% | 4.222% | 0.8223% | -0.7055% | -11.68% | -1.361% | 1.785% | 89.94% | 6.242% | -19.12% | 0.5398% | 0.3432% | -4.376%
| Z-Score (Absolute) |  | 0.02 | 13.55 | 0.67 | 1.06 | 0.49 | 1.17 | 2.8 | 0.2 | 0.02 | 0.9 | 0.09 | 1.12 | 1.59 | 0.68 | 0.79 | 0.22 | 0.04 | 1.89 | 0.15 | 0.38 | 15.25 | 1.13 | 3.14 | 0.17 | 0.14 | 0.66
| Analytical results for Be in OREAS 555b (Certified Value 4.76 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 7.31 | 5 | < 5 | 4.94 | 6.2 | 4.45 | 5.054 | 4.66 | 4.43 | 4.56 | 4.49 | 4.84 | < 5 | 4.4 | 4.8806 | 5.8352 | 4.5 | 4.7425 | 4.34 | 5.1 | 3.7 | 4 | 3.55 | 3.79 | 4.76 | 5 | 4.98 | 1
| 1 | 2 | 6.91 | 5 | < 5 | 4.82 | 5.9 | 4.41 | 5.066 | 4.62 | 4.48 | 4.73 | 4.56 | 4.77 | < 5 | 4.6 | 5.1096 | 5.8646 | 4.6 | 4.7287 | 4.39 | 5.1 | 4.1 | 4 | 3.39 | 4.56 | 4.7 | < 5 | 5 | 3
| 1 | 3 | 7.02 | 5 | < 5 | 4.8 | 6.3 | 4.65 | 5.1058 | 4.72 | 4.56 | 4.71 | 4.65 | 4.82 | < 5 | 4.4 | 4.9104 | 5.8281 | 4.8 | 4.6107 | 4.45 | 4.6 | 3.6 | 4 | 3.33 | 4.16 | 4.73 | 5 | 4.9 | 16
| 1 | 4 | 7.19 | 5 | < 5 | 4.79 | 4.5 | 4.49 | 5.2352 | 4.67 | 4.53 | 4.74 | 4.54 | 4.79 | < 5 | 4.8 | 5.1629 | 5.7588 | 4.7 | 4.7499 | 4.36 | 5.1 | 4 | 4 | 3.29 | 4.26 | 4.73 | < 5 | 5.05 | 4.7558
| 1 | 5 | 7.06 | 5 | < 5 | 4.81 | 5.9 | 5.12 | 5.259 | 4.54 | 4.56 | 4.68 | 4.52 | 4.78 | < 5 | 4.6 | 5.1487 | 5.7739 | 4.9 | 4.7794 | 4.39 | 4.7 | 3.6 | 4 | 3.26 | 5.51 | 4.95 | < 5 | 4.96 | 91
| 1 | 6 | 6.86 | 5 | < 5 | 4.8 | 6 | 4.71 | 5.2369 | 4.59 | 4.43 | 4.65 | 4.65 | 4.82 | < 5 | 4.7 | 5.2162 | 5.6663 | 4.7 | 4.8024 | 4.31 | 5 | 4.1 | 4 | 3.48 | 4.59 | 4.9 | < 5 | 5.05
| Mean |  | 7.0583 | 5 |  | 4.8267 | 5.8 | 4.6383 | 5.1595 | 4.6333 | 4.4983 | 4.6783 | 4.5683 | 4.8033 |  | 4.5833 | 5.0714 | 5.7878 | 4.7 | 4.7356 | 4.3733 | 4.9333 | 3.85 | 4 | 3.3833 | 4.4783 | 4.795 | 5 | 4.99
| Median |  | 7.04 | 5 |  | 4.805 | 5.95 | 4.57 | 5.1705 | 4.64 | 4.505 | 4.695 | 4.55 | 4.805 |  | 4.6 | 5.1292 | 5.801 | 4.7 | 4.7462 | 4.375 | 5.05 | 3.85 | 4 | 3.36 | 4.41 | 4.745 | 5 | 4.99
| Std Dev. |  | 0.1694 | 0 |  | 0.056451 | 0.65727 | 0.26339 | 0.094217 | 0.06377 | 0.06047 | 0.066758 | 0.067355 | 0.027325 |  | 0.16021 | 0.14077 | 0.07149 | 0.14142 | 0.066752 | 0.048442 | 0.22509 | 0.2429 | 0 | 0.11308 | 0.58417 | 0.10368 | 0 | 0.057271
| Rel.Std.Dev. |  | 2.4% | 0% |  | 1.17% | 11.33% | 5.679% | 1.826% | 1.376% | 1.344% | 1.427% | 1.474% | 0.5689% |  | 3.495% | 2.776% | 1.235% | 3.009% | 1.41% | 1.108% | 4.563% | 6.309% | 0% | 3.342% | 13.04% | 2.162% | 0% | 1.148%
| PDM3 |  | 48.42% | 5.135% |  | 1.49% | 21.96% | -2.47% | 8.489% | -2.575% | -5.413% | -1.629% | -3.942% | 0.9998% |  | -3.626% | 6.637% | 21.7% | -1.173% | -0.4243% | -8.042% | 3.733% | -19.05% | -15.89% | -28.86% | -5.834% | 0.8245% | 5.135% | 4.925%
| Z-Score (Absolute) |  | 6.29 | NA | 5.76 | 0.39 | 2.97 | 0.11 | 1.27 | 0.12 | 0.48 | 0 | 0.29 | 0.33 | 5.76 | 0.25 | 1.04 | 2.93 | 0.06 | 0.15 | 0.81 | 0.67 | 2.19 | NA | 3.42 | 0.53 | 0.31 | 3.56 | 0.82
Results from laboratories 2 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 555b (Certified Value 1.48 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 5 | 1.33 | < 5 | 1.56 | 1.53 | 1.57 | 1.5023 | 1.38 | 1.32 | 1.42 | 1.38 | 1.57 | < 5 | 1.4727 | 1.49 | 1.4618 | 1.51 | 1.49 | < 5 | 1.6 | 2.06 | 1.62 | 1.56 | 1.6 | 1.4 |  |  | 1
| 1 | 2 | < 5 | 1.41 | < 5 | 1.55 | 1.52 | 1.59 | 1.4973 | 1.4 | 1.28 | 1.48 | 1.37 | 1.55 | < 5 | 1.406 | 1.41 | 1.4636 | 1.27 | 1.41 | < 5 | 1.5 | 2.04 | 1.56 | 1.54 | 1.5 | 1.42 |  |  | 4
| 1 | 3 | < 5 | 1.28 | < 5 | 1.58 | 1.55 | 1.6 | 1.5093 | 1.38 | 1.32 | 1.6 | 1.45 | 1.58 | < 5 | 1.441 | 1.52 | 1.4444 | 1.34 | 1.5 | < 5 | 1.5 | 2.04 | 1.55 | 1.55 | 1.6 | 1.39 |  |  | 16
| 1 | 4 | < 5 | 1.32 | < 5 | 1.53 | 1.56 | 1.61 | 1.4679 | 1.39 | 1.27 | 1.46 | 1.4 | 1.49 | < 5 | 1.4155 | 1.5 | 1.4531 | 1.45 | 1.47 | < 5 | 1.4 | 2 | 1.58 | 1.52 | 1.6 | 1.41 |  |  | 1.4754
| 1 | 5 | < 5 | 1.31 | < 5 | 1.59 | 1.54 | 1.62 | 1.4909 | 1.4 | 1.29 | 1.53 | 1.34 | 1.6 | < 5 | 1.4746 | 1.52 | 1.4817 | 1.37 | 1.48 | < 5 | 1.4 | 2.03 | 1.73 | 1.51 | 1.6 | 1.42 |  |  | 92
| 1 | 6 | < 5 | 1.28 | < 5 | 1.55 | 1.5 | 1.59 | 1.4794 | 1.39 | 1.28 | 1.36 | 1.42 | 1.49 | < 5 | 1.4244 | 1.49 | 1.4799 | 1.33 | 1.48 | < 5 | 1.5 | 2.04 | 1.76 | 1.56 | 1.5 | 1.45
| Mean |  |  | 1.3217 |  | 1.56 | 1.5333 | 1.5967 | 1.4912 | 1.39 | 1.2933 | 1.475 | 1.3933 | 1.5467 |  | 1.439 | 1.4883 | 1.4641 | 1.3783 | 1.4717 |  | 1.4833 | 2.035 | 1.6333 | 1.54 | 1.5667 | 1.415
| Median |  |  | 1.315 |  | 1.555 | 1.535 | 1.595 | 1.4941 | 1.39 | 1.285 | 1.47 | 1.39 | 1.56 |  | 1.4327 | 1.495 | 1.4627 | 1.355 | 1.48 |  | 1.5 | 2.04 | 1.6 | 1.545 | 1.6 | 1.415
| Std Dev. |  |  | 0.047924 |  | 0.021909 | 0.021602 | 0.017512 | 0.01532 | 0.0089443 | 0.021602 | 0.083845 | 0.038816 | 0.046762 |  | 0.029186 | 0.040702 | 0.014659 | 0.087274 | 0.031885 |  | 0.075277 | 0.019748 | 0.090259 | 0.020976 | 0.05164 | 0.020736
| Rel.Std.Dev. |  |  | 3.626% |  | 1.404% | 1.409% | 1.097% | 1.027% | 0.6435% | 1.67% | 5.684% | 2.786% | 3.023% |  | 2.028% | 2.735% | 1.001% | 6.332% | 2.167% |  | 5.075% | 0.9704% | 5.526% | 1.362% | 3.296% | 1.465%
| PDM3 |  |  | -10.42% |  | 5.731% | 3.923% | 8.216% | 1.067% | -5.791% | -12.34% | -0.03039% | -5.565% | 4.827% |  | -2.468% | 0.8733% | -0.7707% | -6.582% | -0.2563% |  | 0.5344% | 37.92% | 10.7% | 4.375% | 6.182% | -4.097%
| Z-Score (Absolute) |  | 8.93 | 1.5 | 8.93 | 0.61 | 0.37 | 0.93 | 0 | 0.9 | 1.75 | 0.14 | 0.87 | 0.49 | 8.93 | 0.46 | 0.03 | 0.24 | 1 | 0.17 | 8.93 | 0.07 | 4.81 | 1.26 | 0.43 | 0.67 | 0.67
| Analytical results for Ca in OREAS 555b (Certified Value 0.145 wt.%) | [28 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.16 | 0.16 | 0.15 | 0.1447 | 0.16 | 0.13 | 0.14 | 0.14 | 0.14 | 0.15 | 0.15 | 0.1417 | 0.114 | 0.13379 | 0.14279 | 0.151 | 0.14793 | 0.1595 | 0.14 | 0.13 | 0.14 | 0.133 | 0.127 | 0.1415 | 0.15 | 0.14 |  | 1
| 1 | 2 | 0.15 | 0.15 | 0.15 | 0.1402 | 0.15 | 0.13 | 0.14 | 0.14 | 0.15 | 0.15 | 0.15 | 0.1508 | 0.121 | 0.13543 | 0.14521 | 0.151 | 0.14867 | 0.1588 | 0.14 | 0.14 | 0.14 | 0.134 | 0.081 | 0.1394 | 0.15 | 0.14 |  | #N/A
| 1 | 3 | 0.18 | 0.16 | 0.14 | 0.1411 | 0.15 | 0.13 | 0.14 | 0.14 | 0.15 | 0.15 | 0.15 | 0.1483 | 0.114 | 0.14101 | 0.14804 | 0.154 | 0.14729 | 0.1611 | 0.14 | 0.13 | 0.13 | 0.135 | 0.0849 | 0.1371 | 0.15 | 0.14 |  | 16
| 1 | 4 | 0.16 | 0.16 | 0.14 | 0.1399 | 0.11 | 0.13 | 0.14 | 0.14 | 0.15 | 0.15 | 0.15 | 0.1475 | 0.121 | 0.13204 | 0.14949 | 0.153 | 0.14955 | 0.1586 | 0.14 | 0.14 | 0.13 | 0.135 | 0.0885 | 0.1399 | 0.15 | 0.14 |  | 0.14453
| 1 | 5 | 0.15 | 0.16 | 0.14 | 0.1453 | 0.15 | 0.13 | 0.14 | 0.14 | 0.15 | 0.14 | 0.15 | 0.1477 | 0.121 | 0.14162 | 0.14434 | 0.156 | 0.15139 | 0.1578 | 0.13 | 0.13 | 0.14 | 0.135 | 0.0853 | 0.1394 | 0.15 | 0.14 |  | 93
| 1 | 6 | 0.16 | 0.16 | 0.14 | 0.1401 | 0.15 | 0.13 | 0.14 | 0.14 | 0.14 | 0.15 | 0.15 | 0.1796 | 0.114 | 0.13118 | 0.14429 | 0.151 | 0.14884 | 0.1588 | 0.13 | 0.14 | 0.14 | 0.139 | 0.0949 | 0.1375 | 0.15 | 0.14
| Mean |  | 0.16 | 0.15833 | 0.14333 | 0.14188 | 0.145 | 0.13 | 0.14 | 0.14 | 0.14667 | 0.14833 | 0.15 | 0.1526 | 0.1175 | 0.13585 | 0.14569 | 0.15267 | 0.14894 | 0.1591 | 0.13667 | 0.135 | 0.13667 | 0.13517 | 0.0936 | 0.13913 | 0.15 | 0.14
| Median |  | 0.16 | 0.16 | 0.14 | 0.14065 | 0.15 | 0.13 | 0.14 | 0.14 | 0.15 | 0.15 | 0.15 | 0.148 | 0.1175 | 0.13461 | 0.14477 | 0.152 | 0.14876 | 0.1588 | 0.14 | 0.135 | 0.14 | 0.135 | 0.0869 | 0.1394 | 0.15 | 0.14
| Std Dev. |  | 0.010954 | 0.0040825 | 0.005164 | 0.0024564 | 0.017607 | 0 | 0 | 0 | 0.005164 | 0.0040825 | 0 | 0.013561 | 0.0038341 | 0.0044851 | 0.002545 | 0.0020656 | 0.0014307 | 0.0011207 | 0.005164 | 0.0054772 | 0.005164 | 0.0020412 | 0.017009 | 0.0016207 | 0 | 0
| Rel.Std.Dev. |  | 6.847% | 2.578% | 3.603% | 1.731% | 12.14% | 0% | 0% | 0% | 3.521% | 2.752% | 0% | 8.887% | 3.263% | 3.302% | 1.747% | 1.353% | 0.9605% | 0.7044% | 3.779% | 4.057% | 3.779% | 1.51% | 18.17% | 1.165% | 0% | 0%
| PDM3 |  | 10.71% | 9.553% | -0.8255% | -1.829% | 0.3277% | -10.05% | -3.132% | -3.132% | 1.481% | 2.634% | 3.787% | 5.586% | -18.7% | -6.006% | 0.8074% | 5.632% | 3.057% | 10.08% | -5.438% | -6.591% | -5.438% | -6.476% | -35.24% | -3.732% | 3.787% | -3.132%
| Z-Score (Absolute) |  | 1.79 | 1.62 | 0.07 | 0.07 | 0.25 | 1.3 | 0.27 | 0.27 | 0.42 | 0.59 | 0.76 | 1.03 | 2.54 | 0.7 | 0.32 | 1.04 | 0.65 | 1.7 | 0.61 | 0.78 | 0.61 | 0.77 | 5.05 | 0.36 | 0.76 | 0.27
| Analytical results for Cd in OREAS 555b (Certified Value 0.46 ppm) | [28 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | < 1 | 0.51 | < 0.5 | 0.47 | < 0.3 | 0.54 | 0.907 | 0.39 | 0.41 | 0.7 | 0.48 | 0.49 | < 5 | < 1 | 0.45576 | 0.48 | 0.52572 | 0.38 | 0.52 | 1 | 0.6 | 0.34 | 0.36 | 0.5 | < 0.5 | 0.49 |  | 1
| 1 | 2 | < 1 | 0.36 | < 0.5 | 0.45 | 0.5 | 0.49 | 0.88668 | 0.45 | 0.42 | 0.52 | 0.42 | 0.48 | < 5 | < 1 | 0.43985 | 0.48 | 0.54589 | 0.36 | 0.49 | 1 | 0.8 | 0.32 | 0.33 | 0.49 | < 0.5 | 0.5 |  | 5
| 1 | 3 | < 1 | 0.5 | < 0.5 | 0.49 | < 0.3 | 0.52 | 0.91588 | 0.41 | 0.43 | 0.44 | 0.51 | 0.47 | < 5 | < 1 | 0.4754 | 0.47 | 0.5257 | 0.39 | 0.51 | 1 | 0.6 | 0.33 | 0.37 | 0.51 | < 0.5 | 0.49 |  | 16
| 1 | 4 | < 1 | 0.52 | < 0.5 | 0.4 | < 0.3 | 0.52 | 0.93154 | 0.45 | 0.39 | 0.46 | 0.45 | 0.45 | < 5 | < 1 | 0.47782 | 0.49 | 0.52665 | 0.38 | 0.51 | 1 | 0.7 | 0.3 | 0.37 | 0.47 | < 0.5 | 0.49 |  | 0.45646
| 1 | 5 | < 1 | 0.53 | < 0.5 | 0.44 | < 0.3 | 0.51 | 0.88234 | 0.42 | 0.42 | 0.5 | 0.5 | 0.48 | < 5 | < 1 | 0.45274 | 0.46 | 0.50605 | 0.38 | 0.51 | 1 | 0.6 | 0.35 | 0.35 | 0.51 | < 0.5 | 0.47 |  | 94
| 1 | 6 | < 1 | 0.54 | < 0.5 | 0.47 | < 0.3 | 0.51 | 0.93763 | 0.43 | 0.42 | 0.46 | 0.48 | 0.46 | < 5 | < 1 | 0.44056 | 0.46 | 0.50048 | 0.37 | 0.52 | 1 | 0.8 | 0.34 | 0.34 | 0.47 | < 0.5 | 0.51
| Mean |  |  | 0.49333 |  | 0.45333 | 0.5 | 0.515 | 0.91018 | 0.425 | 0.415 | 0.51333 | 0.47333 | 0.47167 |  |  | 0.45702 | 0.47333 | 0.52175 | 0.37667 | 0.51 | 1 | 0.68333 | 0.33 | 0.35333 | 0.49167 |  | 0.49167
| Median |  |  | 0.515 |  | 0.46 | 0.5 | 0.515 | 0.91144 | 0.425 | 0.42 | 0.48 | 0.48 | 0.475 |  |  | 0.45425 | 0.475 | 0.52571 | 0.38 | 0.51 | 1 | 0.65 | 0.335 | 0.355 | 0.495 |  | 0.49
| Std Dev. |  |  | 0.066833 |  | 0.031411 |  | 0.016432 | 0.022714 | 0.023452 | 0.013784 | 0.096056 | 0.033267 | 0.01472 |  |  | 0.016469 | 0.012111 | 0.016352 | 0.010328 | 0.010954 | 0 | 0.098319 | 0.017889 | 0.01633 | 0.018348 |  | 0.013292
| Rel.Std.Dev. |  |  | 13.55% |  | 6.929% |  | 3.191% | 2.496% | 5.518% | 3.321% | 18.71% | 7.028% | 3.121% |  |  | 3.603% | 2.559% | 3.134% | 2.742% | 2.148% | 0% | 14.39% | 5.421% | 4.622% | 3.732% |  | 2.703%
| PDM3 |  |  | 8.079% |  | -0.6843% | 9.539% | 12.83% | 99.4% | -6.892% | -9.082% | 12.46% | 3.697% | 3.332% |  |  | 0.1238% | 3.697% | 14.3% | -17.48% | 11.73% | 119.1% | 49.7% | -27.7% | -22.59% | 7.714% |  | 7.714%
| Z-Score (Absolute) |  | 0.45 | 0.34 | 3.76 | 0.34 | NA | 0.7 | 7.36 | 0.81 | 0.98 | 0.67 | 0 | 0.03 | 34.16 | 0.45 | 0.27 | 0 | 0.82 | 1.63 | 0.62 | NA | NA | 2.42 | 2.02 | 0.31 | 3.76 | 0.31
Results from laboratories 5 and 22 were removed due to their 0.1 ppm reading resolution. Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 555b (Certified Value 148 ppm) | [28 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 0
0
| 1 | 1 | 152.52 | 146.75 | 159 | 149 | 149.4669 | 144 | 143.5 | 150 | 151.5 | 160.5 | 136 | 146 | 144.9142 | 139.2 | 153 | 100 | 113 | 124 | 146 | 152 | 141.19 | [6 empty] | 1
| 1 | 2 | 151.62 | 148.81 | 159 | 149 | 147.9444 | 150.5 | 139.5 | 155 | 146 | 160 | 133 | 144 | 143.7065 | 140.1 | 148 | 100 | 110 | 140 | 150 | 144 | 144.98 | [6 empty] | 6
| 1 | 3 | 151.24 | 147.8 | 155 | 150 | 147.4558 | 150 | 140 | 147.5 | 151.5 | 159 | 134 | 147 | 143.6946 | 144.1 | 150 | 103 | 109 | 123 | 150 | 154 | 149.22 | [6 empty] | 16
| 1 | 4 | 158.72 | 147.55 | 116 | 150 | 149.3181 | 144.5 | 132.5 | 156.5 | 147.5 | 149.5 | 135 | 147 | 147.081 | 138.3 | 150 | 107 | 105 | 135 | 149 | 149 | 148.42 | [6 empty] | 148.4917
| 1 | 5 | 159.28 | 150.93 | 156 | 152 | 148.6684 | 152 | 141.5 | 144.5 | 141.5 | 161 | 137 | 143 | 146.4608 | 140.4 | 150 | 101 | 110 | 164 | 151 | 157 | 140.7 | [6 empty] | 95
| 1 | 6 | 154.68 | 146.4 | 158 | 153 | 149.7606 | 149 | 134.5 | 142.5 | 150 | 154 | 138 | 141 | 147.3449 | 132.9 | 150 | 100 | 108 | 138 | 149 | 155 | 148.68
| Mean |  | 154.6767 | 148.04 | 150.5 | 150.5 | 148.769 | 148.3333 | 138.5833 | 149.3333 | 148 | 157.3333 | 135.5 | 144.6667 | 145.5337 | 139.1667 | 150.1667 | 101.8333 | 109.1667 | 137.3333 | 149.1667 | 151.8333 | 145.5317
| Median |  | 153.6 | 147.675 | 157 | 150 | 148.9932 | 149.5 | 139.75 | 148.75 | 148.75 | 159.5 | 135.5 | 145 | 145.6875 | 139.65 | 150 | 100.5 | 109.5 | 136.5 | 149.5 | 153 | 146.7
| Std Dev. |  | 3.5593 | 1.649 | 16.9794 | 1.6432 | 0.91503 | 3.3116 | 4.2239 | 5.6095 | 3.873 | 4.6007 | 1.8708 | 2.4221 | 1.6516 | 3.6539 | 1.6021 | 2.7869 | 2.6394 | 14.8818 | 1.7224 | 4.7081 | 3.8564
| Rel.Std.Dev. |  | 2.301% | 1.114% | 11.28% | 1.092% | 0.6151% | 2.233% | 3.048% | 3.756% | 2.617% | 2.924% | 1.381% | 1.674% | 1.135% | 2.626% | 1.067% | 2.737% | 2.418% | 10.84% | 1.155% | 3.101% | 2.65%
| PDM3 |  | 4.165% | -0.3042% | 1.352% | 1.352% | 0.1868% | -0.1067% | -6.673% | 0.5668% | -0.3311% | 5.954% | -8.749% | -2.576% | -1.992% | -6.28% | 1.128% | -31.42% | -26.48% | -7.514% | 0.4545% | 2.25% | -1.993%
| Z-Score (Absolute) |  | 1.78 | 0 | 0.66 | 0.66 | 0.2 | 0.08 | 2.54 | 0.35 | 0.01 | 2.5 | 3.37 | 0.91 | 0.67 | 2.39 | 0.57 | 12.42 | 10.45 | 2.88 | 0.3 | 1.02 | 0.67
| Analytical results for Co in OREAS 555b (Certified Value 1.1 wt.%) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*AAS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.1 | 1.08 | 1.1703 | 1.1068 | 1.1 | 0.981 | 1.1427 | 1.055 | 1.05 | 1.065 | 1.09 | 1.105 | 1.0861 | 1.16 | 1.1726 | 1.0664 | 1.06 | 1.1003 | 1.1241 | 1.1 | 1.12 | 1.07 | 1.07 | 1.1553 | 1.0988 | 1.084 |  | 1
| 1 | 2 | 1.1 | 1.05 | 1.1311 | 1.0947 | 1.09 | 1.02 | 1.1162 | 1.065 | 1.06 | 1.07 | 1.105 | 1.075 | 1.1535 | 1.17 | 1.1456 | 1.0295 | 1.01 | 1.1243 | 1.1224 | 1.1 | 1.11 | 1.07 | 1.08 | 1.114 | 1.0879 | 1.105 |  | 36
| 1 | 3 | 1.08 | 1.06 | 1.1806 | 1.0991 | 1.09 | 1.03 | 1.1331 | 1.06 | 1.085 | 1.06 | 1.1 | 1.07 | 1.1406 | 1.21 | 1.1481 | 1.0589 | 1.01 | 1.1112 | 1.1336 | 1.1 | 1.14 | 1.1 | 1.08 | 1.1204 | 1.1339 | 1.099 |  | 16
| 1 | 4 | 1.09 | 1.06 | 1.1629 | 1.1138 | 1.08 | 1.02 | 1.1131 | 1.055 | 1.09 | 1.04 | 1.105 | 1.07 | 1.1325 | 1.18 | 1.1277 | 1.0635 | 1.01 | 1.1154 | 1.1238 | 1.1 | 1.12 | 1.1 | 1.06 | 1.1271 | 1.0712 | 1.058 |  | 1.102
| 1 | 5 | 1.09 | 1.08 | 1.1465 | 1.1562 | 1.09 | 1.04 | 1.1447 | 1.06 | 1.095 | 1.065 | 1.12 | 1.075 | 1.1334 | 1.16 | 1.1636 | 1.0698 | 1.02 | 1.119 | 1.1317 | 1.13 | 1.13 | 1.1 | 1.19 | 1.1026 | 1.1249 | 1.123 |  | 96
| 1 | 6 | 1.09 | 1.08 | 1.1576 | 1.125 | 1.09 | 1.08 | 1.1057 | 1.08 | 1.06 | 1.075 | 1.17 | 1.11 | 1.1349 | 1.12 | 1.1236 | 1.035 | 1.02 | 1.0793 | 1.1296 | 1.11 | 1.13 | 1.09 | 1.22 | 1.1353 | 1.1055 | 1.112
| Mean |  | 1.0917 | 1.0683 | 1.1582 | 1.1159 | 1.09 | 1.0285 | 1.1259 | 1.0625 | 1.0733 | 1.0625 | 1.115 | 1.0842 | 1.1302 | 1.1667 | 1.1469 | 1.0538 | 1.0217 | 1.1083 | 1.1275 | 1.1067 | 1.125 | 1.0883 | 1.1167 | 1.1258 | 1.1037 | 1.0968
| Median |  | 1.09 | 1.07 | 1.1603 | 1.1103 | 1.09 | 1.025 | 1.1246 | 1.06 | 1.0725 | 1.065 | 1.105 | 1.075 | 1.1341 | 1.165 | 1.1468 | 1.0612 | 1.015 | 1.1133 | 1.1269 | 1.1 | 1.125 | 1.095 | 1.08 | 1.1238 | 1.1021 | 1.102
| Std Dev. |  | 0.0075277 | 0.013292 | 0.017561 | 0.022463 | 0.0063246 | 0.03221 | 0.01647 | 0.0093541 | 0.018886 | 0.012145 | 0.028636 | 0.01828 | 0.02295 | 0.029439 | 0.019269 | 0.017198 | 0.019408 | 0.016327 | 0.0046954 | 0.012111 | 0.010488 | 0.01472 | 0.069474 | 0.018261 | 0.023215 | 0.02306
| Rel.Std.Dev. |  | 0.6896% | 1.244% | 1.516% | 2.013% | 0.5802% | 3.132% | 1.463% | 0.8804% | 1.76% | 1.143% | 2.568% | 1.686% | 2.031% | 2.523% | 1.68% | 1.632% | 1.9% | 1.473% | 0.4164% | 1.094% | 0.9323% | 1.352% | 6.222% | 1.622% | 2.103% | 2.102%
| PDM3 |  | -0.9383% | -3.056% | 5.096% | 1.263% | -1.09% | -6.67% | 2.17% | -3.585% | -2.602% | -3.585% | 1.179% | -1.619% | 2.555% | 5.867% | 4.07% | -4.37% | -7.29% | 0.5675% | 2.316% | 0.4228% | 2.086% | -1.241% | 1.33% | 2.159% | 0.1536% | -0.4695%
| Z-Score (Absolute) |  | 0.44 | 1.19 | 1.71 | 0.35 | 0.49 | 2.48 | 0.67 | 1.38 | 1.03 | 1.38 | 0.32 | 0.68 | 0.81 | 1.99 | 1.35 | 1.66 | 2.7 | 0.1 | 0.72 | 0.05 | 0.64 | 0.54 | 0.37 | 0.67 | 0.05 | 0.27
| Analytical results for Cr in OREAS 555b (Certified Value 93 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 97.77 | 114 | 85.1 | 100 | 66 | 91.3 | 105.0416 | 98 | 95 | 99 | 98 | 90 | 83.91 | 107 | 97.4695 | 88 | 94.8525 | 84.7 | 89 | 81 | 87 | 95 | 83 | 95 | 62 | 97 |  | 1
| 1 | 2 | 102 | 97 | 85.8 | 97 | 93 | 88.5 | 106.4986 | 97 | 97 | 102 | 87 | 104 | 82.12 | 109 | 99.413 | 87 | 96.6073 | 85.8 | 87 | 86 | 86 | 94 | 84 | 96 | 63 | 96 |  | 37
| 1 | 3 | 102.1 | 112 | 85.5 | 97 | 70 | 88.7 | 107.6916 | 99 | 98 | 102 | 84 | 94 | 83.26 | 106 | 101.5414 | 87 | 94.6763 | 86.7 | 90 | 82 | 87 | 90 | 85 | 96 | 64 | 108 |  | 16
| 1 | 4 | 101.7 | 70 | 87.9 | 95 | 65 | 88.7 | 107.823 | 99 | 98 | 102 | 94 | 102 | 83.49 | 122 | 101.0181 | 85 | 95.0176 | 84.9 | 86 | 85 | 84 | 89 | 91 | 97 | 66 | 98 |  | 93.2783
| 1 | 5 | 91.41 | 59 | 86.6 | 99 | 79 | 92.3 | 106.2287 | 98 | 100 | 97 | 85 | 103 | 83.74 | 106 | 102.4599 | 87 | 96.1355 | 85.4 | 89 | 81 | 85 | 90 | 89 | 100 | 68 | 96 |  | 97
| 1 | 6 | 101.7 | 112 | 86.8 | 100 | 77 | 89.4 | 112.8133 | 98 | 96 | 95 | 85 | 104 | 83.63 | 107 | 97.3081 | 87 | 95.6294 | 84.7 | 94 | 85 | 85 | 91 | 96 | 97 | 66 | 95
| Mean |  | 99.4467 | 94 | 86.2833 | 98 | 75 | 89.8167 | 107.6828 | 98.1667 | 97.3333 | 99.5 | 88.8333 | 99.5 | 83.3583 | 109.5 | 99.8683 | 86.8333 | 95.4864 | 85.3667 | 89.1667 | 83.3333 | 85.6667 | 91.5 | 88 | 96.8333 | 64.8333 | 98.3333
| Median |  | 101.7 | 104.5 | 86.2 | 98 | 73.5 | 89.05 | 107.0951 | 98 | 97.5 | 100.5 | 86 | 102.5 | 83.56 | 107 | 100.2155 | 87 | 95.3235 | 85.15 | 89 | 83.5 | 85.5 | 90.5 | 87 | 96.5 | 65 | 96.5
| Std Dev. |  | 4.2688 | 23.9082 | 1.0226 | 2 | 10.4881 | 1.598 | 2.7142 | 0.75277 | 1.7512 | 3.0166 | 5.7764 | 5.9917 | 0.64564 | 6.2209 | 2.1612 | 0.98319 | 0.7715 | 0.78401 | 2.7869 | 2.2509 | 1.2111 | 2.429 | 4.98 | 1.7224 | 2.2286 | 4.8442
| Rel.Std.Dev. |  | 4.293% | 25.43% | 1.185% | 2.041% | 13.98% | 1.779% | 2.521% | 0.7668% | 1.799% | 3.032% | 6.503% | 6.022% | 0.7745% | 5.681% | 2.164% | 1.132% | 0.808% | 0.9184% | 3.125% | 2.701% | 1.414% | 2.655% | 5.659% | 1.779% | 3.437% | 4.926%
| PDM3 |  | 6.613% | 0.7737% | -7.499% | 5.062% | -19.6% | -3.711% | 15.44% | 5.241% | 4.347% | 6.67% | -4.765% | 6.67% | -10.63% | 17.39% | 7.065% | -6.909% | 2.367% | -8.482% | -4.408% | -10.66% | -8.16% | -1.906% | -5.659% | 3.811% | -30.49% | 5.419%
| Z-Score (Absolute) |  | 0.73 | 0.14 | 0.7 | 0.57 | 1.93 | 0.32 | 1.63 | 0.59 | 0.5 | 0.74 | 0.43 | 0.74 | 1.02 | 1.82 | 0.78 | 0.64 | 0.3 | 0.8 | 0.39 | 1.03 | 0.77 | 0.14 | 0.52 | 0.44 | 3.04 | 0.61
| Analytical results for Cs in OREAS 555b (Certified Value 2.03 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.5 | 2.12 | 2.13 | 2.21 | 2.0152 | 1.84 | 1.83 | 1.96 | 2.02 | 2.15 | 2.0893 | 2.069 | 2 | 2.0992 | 2.2 | 2 | 2.1 | 2.11 | 1.5 | 2.17 | 1.92 | [6 empty] | 1
| 1 | 2 | 2 | 2.04 | 2.11 | 2.19 | 2.0063 | 1.96 | 1.76 | 1.92 | 1.95 | 2.09 | 2.1301 | 2.1006 | 2.1 | 2.07 | 2 | 1.91 | 2 | 2.03 | 1.6 | 2.12 | 1.95 | [6 empty] | 9
| 1 | 3 | 1.4 | 2.03 | 2.07 | 2.22 | 2.0117 | 1.93 | 1.84 | 1.94 | 2.06 | 2.08 | 2.066 | 2.089 | 2 | 2.0268 | 2.1 | 1.95 | 2 | 2.07 | 1.5 | 2.12 | 1.9 | [6 empty] | 16
| 1 | 4 | 1.5 | 2.03 | 1.48 | 2.19 | 2.0204 | 1.91 | 1.73 | 1.97 | 2.01 | 1.98 | 2.0891 | 2.0466 | 2 | 2.1336 | 2.2 | 1.95 | 2 | 1.97 | 1.6 | 2.12 | 2 | [6 empty] | 2.0321
| 1 | 5 | 1.5 | 2.07 | 2.02 | 2.26 | 2.0234 | 1.97 | 1.84 | 1.92 | 1.94 | 2.12 | 2.1415 | 2.0515 | 2 | 2.1375 | 2.1 | 1.95 | 2 | 1.99 | 1.7 | 2.19 | 1.93 | [6 empty] | 98
| 1 | 6 | 1.5 | 2.15 | 2.04 | 2.27 | 1.9734 | 1.93 | 1.76 | 1.85 | 2.01 | 2.11 | 2.0423 | 2.0607 | 2 | 2.0886 | 2.1 | 1.97 | 1.9 | 2.13 | 1.7 | 2.09 | 2
| Mean |  | 1.5667 | 2.0733 | 1.975 | 2.2233 | 2.0084 | 1.9233 | 1.7933 | 1.9267 | 1.9983 | 2.0883 | 2.093 | 2.0696 | 2.0167 | 2.0926 | 2.1167 | 1.955 | 2 | 2.05 | 1.6 | 2.135 | 1.95
| Median |  | 1.5 | 2.055 | 2.055 | 2.215 | 2.0134 | 1.93 | 1.795 | 1.93 | 2.01 | 2.1 | 2.0892 | 2.0648 | 2 | 2.0939 | 2.1 | 1.95 | 2 | 2.05 | 1.6 | 2.12 | 1.94
| Std Dev. |  | 0.21602 | 0.050859 | 0.24599 | 0.034448 | 0.018181 | 0.046332 | 0.048854 | 0.04274 | 0.04535 | 0.058452 | 0.037591 | 0.02132 | 0.040825 | 0.041478 | 0.075277 | 0.029496 | 0.063246 | 0.064498 | 0.089443 | 0.037283 | 0.041952
| Rel.Std.Dev. |  | 13.79% | 2.453% | 12.46% | 1.549% | 0.9052% | 2.409% | 2.724% | 2.218% | 2.269% | 2.799% | 1.796% | 1.03% | 2.024% | 1.982% | 3.556% | 1.509% | 3.162% | 3.146% | 5.59% | 1.746% | 2.151%
| PDM3 |  | -22.9% | 2.031% | -2.808% | 9.412% | -1.165% | -5.351% | -11.75% | -5.187% | -1.66% | 2.769% | 3.001% | 1.845% | -0.7579% | 2.98% | 4.163% | -3.793% | -1.578% | 0.8825% | -21.26% | 5.065% | -4.039%
| Z-Score (Absolute) |  | 3.73 | 0.55 | 0.28 | 1.81 | 0 | 0.72 | 1.81 | 0.69 | 0.08 | 0.67 | 0.71 | 0.52 | 0.07 | 0.71 | 0.91 | 0.45 | 0.07 | 0.35 | 3.44 | 1.07 | 0.49
| Analytical results for Cu in OREAS 555b (Certified Value 4.58 wt.%) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.48 | 4.598 | 4.6528 | 4.4466 | 4.65 | 4.48 | 4.6825 | 4.51 | 4.56 | 4.45 | 4.59 | 4.54 | 4.6324 | 4.66 | 4.3696 | 4.1247 | 4.43 | 4.7242 | 1.55 | 4.72 | 4.697 | 4.68 | > 1 | 4.232 | 4.3612 | 4.4452 | 4.546 | 1
| 1 | 2 | 4.44 | 4.623 | 4.615 | 4.2842 | 4.6 | 4.3 | 4.6826 | 4.55 | 4.57 | 4.48 | 4.59 | 4.48 | 4.8255 | 4.64 | 4.6414 | 4.1614 | 4.43 | 4.711 | 1.5689 | 4.64 | 4.651 | 4.72 | > 1 | 4.958 | 4.1439 | 4.5119 | 4.724 | 39
| 1 | 3 | 4.4 | 4.545 | 4.665 | 4.4508 | 4.63 | 4.42 | 4.8605 | 4.52 | 4.63 | 4.56 | 4.64 | 4.57 | 4.8214 | 4.7 | 4.4302 | 4.1939 | 4.49 | 4.7048 | 1.5378 | 4.66 | 4.761 | 4.74 | > 1 | 4.328 | 4.2066 | 4.5355 | 4.629 | 16
| 1 | 4 | 4.5 | 4.64 | 4.68 | 4.4134 | 4.57 | 4.34 | 4.6215 | 4.51 | 4.61 | 4.41 | 4.58 | 4.57 | 4.8206 | 4.73 | 4.3368 | 4.1698 | 4.49 | 4.8017 | 1.5427 | 4.7 | 4.681 | 4.7 | > 1 | 4.544 | 4.4252 | 4.4723 | 4.543 | 4.5807
| 1 | 5 | 4.44 | 4.616 | 4.655 | 4.5462 | 4.62 | 4.44 | 4.8784 | 4.49 | 4.59 | 4.44 | 4.57 | 4.54 | 4.8115 | 4.68 | 4.4548 | 4.1389 | 4.56 | 4.6865 | 1.3745 | 4.72 | 4.61 | 4.79 | > 1 | 4.921 | 4.2808 | 4.5559 | 4.649 | 99
| 1 | 6 | 4.48 | 4.529 | 4.65 | 4.5307 | 4.63 | 4.35 | 4.7499 | 4.49 | 4.52 | 4.57 | 4.58 | 4.57 | 4.8223 | 4.66 | 4.6856 | 4.1147 | 4.53 | 4.6457 | 1.6455 | 4.7 | 4.71 | 4.8 | > 1 | 4.645 | 4.4047 | 4.53 | 4.568
| Mean |  | 4.4567 | 4.5918 | 4.653 | 4.4453 | 4.6167 | 4.3883 | 4.7459 | 4.5117 | 4.58 | 4.485 | 4.5917 | 4.545 | 4.7889 | 4.6783 | 4.4864 | 4.1506 | 4.4883 | 4.7123 | 1.5366 | 4.69 | 4.685 | 4.7383 |  | 4.6047 | 4.3037 | 4.5085 | 4.6098
| Median |  | 4.46 | 4.607 | 4.6539 | 4.4487 | 4.625 | 4.385 | 4.7163 | 4.51 | 4.58 | 4.465 | 4.585 | 4.555 | 4.821 | 4.67 | 4.4425 | 4.1502 | 4.49 | 4.7079 | 1.5464 | 4.7 | 4.689 | 4.73 |  | 4.5945 | 4.321 | 4.521 | 4.5985
| Std Dev. |  | 0.036697 | 0.044844 | 0.021587 | 0.094315 | 0.028048 | 0.068823 | 0.10413 | 0.022286 | 0.038987 | 0.065955 | 0.024833 | 0.035071 | 0.076837 | 0.032506 | 0.14416 | 0.029855 | 0.052313 | 0.05158 | 0.088756 | 0.032863 | 0.051657 | 0.048339 |  | 0.29854 | 0.11292 | 0.041863 | 0.071031
| Rel.Std.Dev. |  | 0.8234% | 0.9766% | 0.4639% | 2.122% | 0.6075% | 1.568% | 2.194% | 0.494% | 0.8512% | 1.471% | 0.5408% | 0.7716% | 1.604% | 0.6948% | 3.213% | 0.7193% | 1.166% | 1.095% | 5.776% | 0.7007% | 1.103% | 1.02% |  | 6.484% | 2.624% | 0.9285% | 1.541%
| PDM3 |  | -2.708% | 0.2425% | 1.577% | -2.956% | 0.7846% | -4.2% | 3.606% | -1.508% | -0.01581% | -2.09% | 0.2389% | -0.7799% | 4.546% | 2.131% | -2.059% | -9.391% | -2.017% | 2.873% | -66.46% | 2.386% | 2.276% | 3.441% |  | 0.5227% | -6.047% | -1.577% | 0.6355%
| Z-Score (Absolute) |  | 0.88 | 0.04 | 0.46 | 0.95 | 0.21 | 1.34 | 1.09 | 0.5 | 0.04 | 0.68 | 0.04 | 0.28 | 1.38 | 0.63 | 0.68 | 2.96 | 0.66 | 0.86 | 20.71 | 0.71 | 0.67 | 1.04 | NA | 0.13 | 1.92 | 0.53 | 0.16
| Analytical results for Dy in OREAS 555b (Certified Value 4.68 ppm) | [28 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 3.5 | 4.35 | 5.5 | 6.7001 | 5 | 5.1324 | 4.7043 | 4.29 | 4.43 | 3.85 | 5.03 | [16 empty] | 1
| 1 | 2 | 3.6 | 4.6 | 5.2 | 6.5535 | 4.8 | 5.1621 | 4.5993 | 3.99 | 4.42 | 4.57 | 5.13 | [16 empty] | 10
| 1 | 3 | 3.4 | 4.54 | 5.6 | 6.4503 | 5.1 | 5.7885 | 4.5824 | 4.16 | 4.34 | 3.92 | 5.03 | [16 empty] | 16
| 1 | 4 | 3.4 | 4.62 | 4.3 | 6.3918 | 5.4 | 4.9802 | 4.6949 | 4.41 | 4.28 | 4.26 | 4.87 | [16 empty] | 4.6792
| 1 | 5 | 3.4 | 4.73 | 5.7 | 6.3334 | 5 | 5.9345 | 4.7296 | 4.24 | 4.39 | 4.65 | 5.06 | [16 empty] | 100
| 1 | 6 | 3.3 | 4.71 | 5.9 | 6.7112 | 5.2 | 5.81 | 4.7724 | 4.56 | 4.56 | 4.36 | 4.93
| Mean |  | 3.4333 | 4.5917 | 5.3667 | 6.5234 | 5.0833 | 5.4679 | 4.6805 | 4.275 | 4.4033 | 4.2683 | 5.0083
| Median |  | 3.4 | 4.61 | 5.55 | 6.5019 | 5.05 | 5.4753 | 4.6996 | 4.265 | 4.405 | 4.31 | 5.03
| Std Dev. |  | 0.10328 | 0.1379 | 0.57155 | 0.15883 | 0.20412 | 0.41985 | 0.074601 | 0.19766 | 0.094798 | 0.32908 | 0.09347
| Rel.Std.Dev. |  | 3.008% | 3.003% | 10.65% | 2.435% | 4.016% | 7.678% | 1.594% | 4.624% | 2.153% | 7.71% | 1.866%
| PDM3 |  | -26.63% | -1.87% | 14.69% | 39.41% | 8.637% | 16.86% | 0.02801% | -8.638% | -5.895% | -8.78% | 7.035%
| Z-Score (Absolute) |  | 2.07 | 0.15 | 1.14 | 3.06 | 0.67 | 1.31 | 0 | 0.67 | 0.46 | 0.69 | 0.55
| Analytical results for Er in OREAS 555b (Certified Value 2.76 ppm) | [28 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 2.2 | 2.63 | 3.5 | 3.7274 | 0.8 | 3.4859 | 2.9062 | 2.28 | 2.66 | 2.3 | 3.04 | [16 empty] | 1
| 1 | 2 | 2.3 | 2.68 | 3.2 | 3.6556 | < 0.2 | 3.3115 | 2.7987 | 2.09 | 2.62 | 2.65 | 3.05 | [16 empty] | 11
| 1 | 3 | 2.1 | 2.71 | 3.4 | 3.7023 | 0.6 | 3.909 | 2.918 | 2.17 | 2.56 | 2.34 | 3.05 | [16 empty] | 16
| 1 | 4 | 2.1 | 2.73 | 2.5 | 3.6584 | 0.2 | 3.2341 | 2.885 | 2.37 | 2.52 | 2.5 | 2.93 | [16 empty] | 2.7576
| 1 | 5 | 2.1 | 2.7 | 3.3 | 3.5305 | 0.3 | 3.3326 | 2.9293 | 2.22 | 2.54 | 2.9 | 3.05 | [16 empty] | 101
| 1 | 6 | 2.1 | 2.81 | 3.5 | 3.7352 | < 0.2 | 3.2951 | 2.9199 | 2.16 | 2.75 | 2.55 | 2.91
| Mean |  | 2.15 | 2.71 | 3.2333 | 3.6682 | 0.475 | 3.428 | 2.8928 | 2.215 | 2.6083 | 2.54 | 3.005
| Median |  | 2.1 | 2.705 | 3.35 | 3.6803 | 0.45 | 3.3221 | 2.9121 | 2.195 | 2.59 | 2.525 | 3.045
| Std Dev. |  | 0.083666 | 0.059666 | 0.37771 | 0.075325 | 0.27538 | 0.25005 | 0.048565 | 0.098944 | 0.086814 | 0.21954 | 0.066257
| Rel.Std.Dev. |  | 3.891% | 2.202% | 11.68% | 2.053% | 57.97% | 7.294% | 1.679% | 4.467% | 3.328% | 8.644% | 2.205%
| PDM3 |  | -22.03% | -1.725% | 17.25% | 33.02% | -82.77% | 24.31% | 4.905% | -19.68% | -5.412% | -7.89% | 8.973%
| Z-Score (Absolute) |  | NA | 0.27 | NA | 2.51 | NA | 1.82 | 0.27 | 1.7 | 0.56 | 0.76 | 0.59
Results from laboratories 2, 5 and 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 555b (Certified Value 0.91 ppm) | [28 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 18 | 19 | 23 | 25 | 26 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [17 empty] | 2
3
| 1 | 1 | 0.6 | 0.86 | 0.93 | 0.93603 | 1.2 | 0.90075 | 0.93 | 0.85 | 0.72 | 0.96 | [17 empty] | 1
| 1 | 2 | 0.6 | 0.91 | 0.93 | 0.92054 | 1.3 | 0.91029 | 0.87 | 0.83 | 0.83 | 0.98 | [17 empty] | 12
| 1 | 3 | 0.6 | 0.9 | 0.92 | 0.93314 | 1.2 | 0.87028 | 0.9 | 0.83 | 0.71 | 0.93 | [17 empty] | 16
| 1 | 4 | 0.6 | 0.89 | 0.7 | 0.9739 | 1.3 | 0.92428 | 0.96 | 0.8 | 0.73 | 0.95 | [17 empty] | 0.91493
| 1 | 5 | 0.6 | 0.92 | 0.95 | 0.9456 | 1.2 | 0.91974 | 0.92 | 0.83 | 0.84 | 0.98 | [17 empty] | 102
| 1 | 6 | 0.5 | 0.96 | 0.96 | 0.9412 | 1.2 | 0.91347 | 0.91 | 0.86 | 0.79 | 0.98
| Mean |  | 0.58333 | 0.90667 | 0.89833 | 0.94173 | 1.2333 | 0.90647 | 0.915 | 0.83333 | 0.77 | 0.96333
| Median |  | 0.6 | 0.905 | 0.93 | 0.93861 | 1.2 | 0.91188 | 0.915 | 0.83 | 0.76 | 0.97
| Std Dev. |  | 0.040825 | 0.033267 | 0.098268 | 0.017919 | 0.05164 | 0.01949 | 0.030166 | 0.020656 | 0.057619 | 0.020656
| Rel.Std.Dev. |  | 6.999% | 3.669% | 10.94% | 1.903% | 4.187% | 2.15% | 3.297% | 2.479% | 7.483% | 2.144%
| PDM3 |  | -36.24% | -0.9036% | -1.814% | 2.929% | 34.8% | -0.9254% | 0.007251% | -8.919% | -15.84% | 5.29%
| Z-Score (Absolute) |  | NA | 0 | 0.25 | 1.09 | NA | 0 | 0.26 | 2.27 | 4.22 | 1.76
Results from laboratories 2 and 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 555b (Certified Value 3.83 wt.%) | [28 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.76 | 3.96 | 3.72 | 3.94 | 4.05 | 3.44 | 3.74 | 3.53 | 3.68 | 3.7 | 3.91 | 3.8766 | 3.847 | 3.815 | 3.7467 | 3.92 | 3.8118 | 3.8334 | 3.77 | 3.71 | 3.82 | 4.03 | 3.2 | 4.0091 | 3.93 | 3.9 |  | 1
| 1 | 2 | 3.7 | 3.72 | 3.79 | 3.88 | 4.05 | 3.4 | 3.7 | 3.57 | 3.84 | 3.76 | 3.86 | 4.0118 | 3.791 | 3.9796 | 3.7567 | 3.93 | 3.8868 | 3.8998 | 3.74 | 3.82 | 3.82 | 4.08 | 3.24 | 3.9453 | 3.92 | 3.9 |  | #N/A
| 1 | 3 | 3.66 | 3.92 | 3.71 | 3.89 | 3.98 | 3.5 | 3.74 | 3.61 | 3.85 | 3.81 | 3.88 | 3.9254 | 3.903 | 3.9065 | 3.7214 | 4.02 | 3.812 | 3.9462 | 3.72 | 3.7 | 3.85 | 4.17 | 3.39 | 3.982 | 3.91 | 3.82 |  | 16
| 1 | 4 | 3.75 | 4.08 | 3.72 | 3.83 | 2.96 | 3.51 | 3.71 | 3.58 | 3.83 | 3.74 | 3.86 | 3.9637 | 3.903 | 3.8078 | 3.7708 | 3.94 | 3.8816 | 3.9203 | 3.73 | 3.77 | 3.77 | 4 | 3.58 | 3.9587 | 3.89 | 3.93 |  | 3.8319
| 1 | 5 | 3.61 | 4.06 | 3.74 | 4.02 | 3.94 | 3.52 | 3.7 | 3.61 | 3.79 | 3.72 | 3.9 | 3.9754 | 3.833 | 3.9702 | 3.7168 | 4.04 | 3.8699 | 3.8671 | 3.64 | 3.69 | 3.83 | 4.25 | 3.41 | 3.9862 | 3.95 | 3.92 |  | 103
| 1 | 6 | 3.65 | 3.91 | 3.72 | 3.87 | 3.98 | 3.55 | 3.67 | 3.55 | 3.76 | 3.83 | 3.89 | 3.9856 | 3.826 | 3.7954 | 3.767 | 3.99 | 3.9039 | 3.8883 | 3.64 | 3.82 | 3.87 | 4.03 | 3.53 | 3.9641 | 3.92 | 3.88
| Mean |  | 3.6883 | 3.9417 | 3.7333 | 3.905 | 3.8267 | 3.4867 | 3.71 | 3.575 | 3.7917 | 3.76 | 3.8833 | 3.9564 | 3.8505 | 3.8791 | 3.7466 | 3.9733 | 3.861 | 3.8925 | 3.7067 | 3.7517 | 3.8267 | 4.0933 | 3.3917 | 3.9742 | 3.92 | 3.8917
| Median |  | 3.68 | 3.94 | 3.72 | 3.885 | 3.98 | 3.505 | 3.705 | 3.575 | 3.81 | 3.75 | 3.885 | 3.9695 | 3.84 | 3.8608 | 3.7517 | 3.965 | 3.8757 | 3.8941 | 3.725 | 3.74 | 3.825 | 4.055 | 3.4 | 3.9731 | 3.92 | 3.9
| Std Dev. |  | 0.059133 | 0.12968 | 0.029439 | 0.066558 | 0.42679 | 0.055737 | 0.026833 | 0.032094 | 0.064317 | 0.05099 | 0.020656 | 0.048301 | 0.044658 | 0.084098 | 0.022921 | 0.050465 | 0.03958 | 0.039662 | 0.054283 | 0.059805 | 0.033862 | 0.097297 | 0.15145 | 0.022805 | 0.02 | 0.0392
| Rel.Std.Dev. |  | 1.603% | 3.29% | 0.7886% | 1.704% | 11.15% | 1.599% | 0.7233% | 0.8977% | 1.696% | 1.356% | 0.5319% | 1.221% | 1.16% | 2.168% | 0.6118% | 1.27% | 1.025% | 1.019% | 1.464% | 1.594% | 0.8849% | 2.377% | 4.465% | 0.5738% | 0.5102% | 1.007%
| PDM3 |  | -3.748% | 2.863% | -2.573% | 1.906% | -0.1377% | -9.011% | -3.182% | -6.705% | -1.051% | -1.877% | 1.341% | 3.248% | 0.4842% | 1.23% | -2.228% | 3.69% | 0.7581% | 1.581% | -3.269% | -2.095% | -0.1377% | 6.821% | -11.49% | 3.713% | 2.298% | 1.559%
| Z-Score (Absolute) |  | 1.13 | 0.78 | 0.79 | 0.5 | 0.09 | 2.65 | 0.97 | 1.99 | 0.35 | 0.59 | 0.34 | 0.89 | 0.09 | 0.31 | 0.69 | 1.02 | 0.17 | 0.41 | 0.99 | 0.66 | 0.09 | 1.92 | 3.37 | 1.02 | 0.61 | 0.4
| Analytical results for Ga in OREAS 555b (Certified Value 28.4 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 24.7 | 27.1 | 29.2 | 28.9 | 28.3 | 27.4788 | 27.9 | 25.7 | 29.6 | 29.1 | 28 | 26.54 | 4 | 29.5249 | 27.6 | 28.3614 | 29.29 | 30.42 | 29 | 31 | 23.7 | 28.94 | 26.63 |  |  |  |  | 1
| 1 | 2 | 23.91 | 27.6 | 29.05 | 29.3 | 29.4 | 27.7981 | 29.7 | 25.4 | 30.6 | 28.7 | 27.8 | 25.51 | 5 | 28.5861 | 27.7 | 28.2802 | 29.43 | 29.09 | 27 | 31 | 23.3 | 28.76 | 26.95 |  |  |  |  | 13
| 1 | 3 | 25.68 | 26.6 | 29 | 29.3 | 30.1 | 27.7421 | 29.6 | 25.5 | 29.6 | 29.3 | 27.9 | 26.12 | 3 | 29.7127 | 28.7 | 27.8481 | 30.27 | 30.27 | 28 | 31 | 22.8 | 29.11 | 26.25 |  |  |  |  | 16
| 1 | 4 | 25.17 | 27.3 | 29.55 | 21.5 | 29.2 | 27.73 | 27.9 | 24.6 | 30.6 | 28.9 | 26.4 | 26.24 | 3 | 29.8104 | 27.6 | 28.8325 | 29.43 | 29.6 | 29 | 31 | 23.5 | 28.37 | 27.8 |  |  |  |  | 28.3601
| 1 | 5 | 24.89 | 28.1 | 29.52 | 27.8 | 29.9 | 27.8364 | 29 | 26.2 | 29 | 27.5 | 28 | 26.37 | < 1 | 30.3144 | 27.3 | 28.9056 | 29.49 | 30.05 | 27 | 31 | 26.6 | 28.56 | 26.84 |  |  |  |  | 104
| 1 | 6 | 26.02 | 26.8 | 28.85 | 28.4 | 30.3 | 27.3838 | 29.3 | 25.8 | 30.1 | 29 | 27.5 | 26.38 | 6 | 28.7181 | 27.6 | 28.4022 | 29.03 | 30.16 | 28 | 32 | 25.8 | 28.87 | 26.91
| Mean |  | 25.0617 | 27.25 | 29.195 | 27.5333 | 29.5333 | 27.6616 | 28.9 | 25.5333 | 29.9167 | 28.75 | 27.6 | 26.1933 | 4.2 | 29.4444 | 27.75 | 28.4383 | 29.49 | 29.9317 | 28 | 31.1667 | 24.2833 | 28.7683 | 26.8967
| Median |  | 25.03 | 27.2 | 29.125 | 28.65 | 29.65 | 27.736 | 29.15 | 25.6 | 29.85 | 28.95 | 27.85 | 26.305 | 4 | 29.6188 | 27.6 | 28.3818 | 29.43 | 30.105 | 28 | 31 | 23.6 | 28.815 | 26.875
| Std Dev. |  | 0.74821 | 0.54681 | 0.28627 | 3.0111 | 0.73394 | 0.18489 | 0.8124 | 0.53541 | 0.63377 | 0.6442 | 0.61644 | 0.36352 | 1.3038 | 0.66837 | 0.48477 | 0.38838 | 0.41646 | 0.49749 | 0.89443 | 0.40825 | 1.5355 | 0.26784 | 0.51216
| Rel.Std.Dev. |  | 2.985% | 2.007% | 0.9805% | 10.94% | 2.485% | 0.6684% | 2.811% | 2.097% | 2.118% | 2.241% | 2.233% | 1.388% | 31.04% | 2.27% | 1.747% | 1.366% | 1.412% | 1.662% | 3.194% | 1.31% | 6.323% | 0.931% | 1.904%
| PDM3 |  | -11.63% | -3.914% | 2.944% | -2.915% | 4.137% | -2.463% | 1.904% | -9.968% | 5.488% | 1.375% | -2.68% | -7.64% | -85.19% | 3.823% | -2.151% | 0.2757% | 3.984% | 5.541% | -1.27% | 9.896% | -14.38% | 1.439% | -5.16%
| Z-Score (Absolute) |  | 1.78 | 0.5 | 0.65 | 0.33 | 0.84 | 0.25 | 0.47 | 1.5 | 1.07 | 0.38 | 0.29 | 1.12 | NA | 0.79 | 0.2 | 0.2 | 0.82 | 1.08 | NA | NA | 2.24 | 0.4 | 0.7
Results from laboratories 14, 21 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 555b (Certified Value 5.18 ppm) | [28 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 3.9 | 4.78 | 6.3 | 6.644 | 3 | 12.6157 | 6.1562 | 4.12 | 4.88 | 4.37 | 5.24 | [16 empty] | 1
| 1 | 2 | 4.4 | 4.87 | 6.1 | 6.6036 | 3 | 12.7878 | 5.8956 | 4.2 | 4.84 | 5.06 | 5.29 | [16 empty] | 14
| 1 | 3 | 3.7 | 4.84 | 6.5 | 6.4406 | 4 | 11.9916 | 5.8718 | 4.19 | 4.73 | 4.34 | 5.38 | [16 empty] | 16
| 1 | 4 | 3.8 | 4.97 | 4.9 | 6.4035 | 4 | 12.1396 | 6.0456 | 4.04 | 4.56 | 4.8 | 5.35 | [16 empty] | 5.1774
| 1 | 5 | 3.8 | 5 | 6.4 | 6.3392 | 4 | 12.8721 | 6.1879 | 4.04 | 4.75 | 5.4 | 5.32 | [16 empty] | 105
| 1 | 6 | 3.7 | 5.06 | 6.5 | 6.5433 | 3 | 12.7949 | 6.0803 | 4.03 | 4.96 | 4.94 | 5.18
| Mean |  | 3.8833 | 4.92 | 6.1167 | 6.4957 | 3.5 | 12.5336 | 6.0396 | 4.1033 | 4.7867 | 4.8183 | 5.2933
| Median |  | 3.8 | 4.92 | 6.35 | 6.492 | 3.5 | 12.7018 | 6.0629 | 4.08 | 4.795 | 4.87 | 5.305
| Std Dev. |  | 0.26394 | 0.10677 | 0.61455 | 0.11994 | 0.54772 | 0.37503 | 0.13128 | 0.078145 | 0.13967 | 0.41029 | 0.073666
| Rel.Std.Dev. |  | 6.797% | 2.17% | 10.05% | 1.847% | 15.65% | 2.992% | 2.174% | 1.904% | 2.918% | 8.515% | 1.392%
| PDM3 |  | -25% | -4.972% | 18.14% | 25.46% | -32.4% | 142.1% | 16.65% | -20.75% | -7.548% | -6.936% | 2.239%
| Z-Score (Absolute) |  | 0.85 | 0.13 | 0.7 | 0.97 | NA | 5.17 | 0.65 | 0.7 | 0.22 | 0.2 | 0.13
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 555b (Certified Value 0.2 ppm) | [28 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [10 empty] | 2
2
| 1 | 1 | 0.11 | 2 | < 0.1 | 0.54 | 0.2 | 0.18 | 0.22 | 0.2 | 0.21 | 1.0846 | 1.8 | 0.20922 | NA | 6.7 | 1.69 | 1.7 | 0.09 | [10 empty] | 1
| 1 | 2 | 0.14 | 2.2 | < 0.1 | 0.47 | 0.2 | 0.17 | 0.22 | 0.25 | 0.25 | 1.0325 | 1.7 | 0.20414 | NA | 6.1 | 1.66 | 1.6 | 0.1 | [10 empty] | 15
| 1 | 3 | 0.08 | 2 | < 0.1 | 0.61 | 0.21 | 0.17 | 0.21 | 0.21 | 0.23 | 1.0929 | 1.9 | 0.20967 | NA | 7.6 | 1.6 | 1.6 | 0.09 | [10 empty] | 16
| 1 | 4 | < 0.05 | 1.9 | < 0.1 | 0.49 | 0.19 | 0.17 | 0.21 | 0.19 | 0.21 | 1.0282 | 1.9 | 0.22287 | NA | 7 | 1.7 | 1.6 | 0.1 | [10 empty] | 0.19654
| 1 | 5 | < 0.05 | 2 | < 0.1 | 0.49 | 0.21 | 0.18 | 0.21 | 0.17 | 0.28 | 1.0953 | 1.8 | 0.22477 | NA | 6.3 | 1.9 | 1.7 | 0.1 | [10 empty] | 106
| 1 | 6 | < 0.05 | 2.2 | < 0.1 | 0.45 | 0.2 | 0.17 | 0.2 | 0.23 | 0.26 | 1.011 | 1.9 | 0.224 | NA | 5.9 | 1.85 | 1.6 | 0.1
| Mean |  | 0.11 | 2.05 |  | 0.50833 | 0.20167 | 0.17333 | 0.21167 | 0.20833 | 0.24 | 1.0574 | 1.8333 | 0.21578 |  | 6.6 | 1.7333 | 1.6333 | 0.096667
| Median |  | 0.11 | 2 |  | 0.49 | 0.2 | 0.17 | 0.21 | 0.205 | 0.24 | 1.0586 | 1.85 | 0.21627 |  | 6.5 | 1.695 | 1.6 | 0.1
| Std Dev. |  | 0.03 | 0.12247 |  | 0.058109 | 0.0075277 | 0.005164 | 0.0075277 | 0.028577 | 0.028284 | 0.037577 | 0.08165 | 0.0091038 |  | 0.63246 | 0.11622 | 0.05164 | 0.005164
| Rel.Std.Dev. |  | 27.27% | 5.974% |  | 11.43% | 3.733% | 2.979% | 3.556% | 13.72% | 11.79% | 3.554% | 4.454% | 4.219% |  | 9.583% | 6.705% | 3.162% | 5.342%
| PDM3 |  | -44.03% | 943% |  | 158.6% | 2.609% | -11.81% | 7.697% | 6.001% | 22.11% | 438% | 832.8% | 9.789% |  | 3258% | 781.9% | 731% | -50.82%
| Z-Score (Absolute) |  | 1.28 | NA | 1.44 | 2.68 | 0.07 | 0.33 | 0.01 | 0.01 | 0.27 | 7.62 | NA | 0.05 | NA | NA | 13.7 | NA | 1.02
Results from laboratories 4, 17, 22 and 26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 555b (Certified Value 5.37 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 5.78 | 5.36 | 5.5 | 5.62 | 6.9174 | 5 | 5.1 | 5.1 | 5.7 | 5.5 | 11 | 5.4689 | 5.8188 | 5.2 | 5.4269 | 5.46 | 5.23 | 5.4 | 4.75 | 5 | 5.39 | 5.5 | 5.3 |  |  |  |  | 1
| 1 | 2 | 5.27 | 5.3 | 5.4 | 5.85 | 6.8316 | 5 | 5.1 | 5.5 | 5.5 | 5.4 | 8 | 5.3938 | 5.5672 | 5.3 | 5.2444 | 5.07 | 5.13 | 5.1 | 4.58 | 5.77 | 5.26 | 5.2 | 5.4 |  |  |  |  | 19
| 1 | 3 | 5.46 | 5.37 | 5.3 | 5.71 | 6.735 | 5.1 | 5.3 | 5.3 | 5.7 | 5.5 | 7 | 5.5771 | 5.6511 | 5.3 | 5.4682 | 5.22 | 5.24 | 4.9 | 4.92 | 4.97 | 5.45 | 5.5 | 5.4 |  |  |  |  | 16
| 1 | 4 | 5.54 | 5.23 | 4 | 5.78 | 6.7271 | 5 | 4.9 | 5.3 | 5.5 | 5.3 | 11 | 5.4905 | 5.4021 | 5.2 | 5.3397 | 5.6 | 5.23 | 5 | 4.45 | 5.43 | 5.32 | 5.3 | 5.6 |  |  |  |  | 5.3705
| 1 | 5 | 5.62 | 5.35 | 5.3 | 5.8 | 6.7522 | 5.1 | 5.2 | 5.4 | 5.2 | 5.5 | 10 | 5.6145 | 5.3897 | 5.2 | 5.3613 | 5.34 | 5.3 | 5.1 | 4.76 | 5.88 | 5.36 | 5.6 | 5.5 |  |  |  |  | 107
| 1 | 6 | 5.64 | 5.53 | 5.4 | 5.71 | 6.9336 | 5 | 5 | 5.3 | 5.6 | 5.2 | 6 | 5.3924 | 5.4259 | 5.3 | 5.4541 | 5.07 | 5.28 | 5.2 | 4.87 | 5.55 | 5.52 | 5.4 | 5.5
| Mean |  | 5.5517 | 5.3567 | 5.15 | 5.745 | 6.8162 | 5.0333 | 5.1 | 5.3167 | 5.5333 | 5.4 | 8.8333 | 5.4895 | 5.5425 | 5.25 | 5.3824 | 5.2933 | 5.235 | 5.1167 | 4.7217 | 5.4333 | 5.3833 | 5.4167 | 5.45
| Median |  | 5.58 | 5.355 | 5.35 | 5.745 | 6.7919 | 5 | 5.1 | 5.3 | 5.55 | 5.45 | 9 | 5.4797 | 5.4965 | 5.25 | 5.3941 | 5.28 | 5.235 | 5.1 | 4.755 | 5.49 | 5.375 | 5.45 | 5.45
| Std Dev. |  | 0.17463 | 0.099532 | 0.56833 | 0.08167 | 0.092617 | 0.05164 | 0.14142 | 0.13292 | 0.18619 | 0.12649 | 2.137 | 0.091976 | 0.17054 | 0.054772 | 0.084595 | 0.21407 | 0.058907 | 0.17224 | 0.17747 | 0.38182 | 0.092664 | 0.1472 | 0.10488
| Rel.Std.Dev. |  | 3.146% | 1.858% | 11.04% | 1.422% | 1.359% | 1.026% | 2.773% | 2.5% | 3.365% | 2.342% | 24.19% | 1.675% | 3.077% | 1.043% | 1.572% | 4.044% | 1.125% | 3.366% | 3.759% | 7.027% | 1.721% | 2.717% | 1.924%
| PDM3 |  | 3.374% | -0.2571% | -4.105% | 6.974% | 26.92% | -6.278% | -5.036% | -1.002% | 3.033% | 0.5498% | 64.48% | 2.217% | 3.203% | -2.243% | 0.223% | -1.436% | -2.523% | -4.726% | -12.08% | 1.171% | 0.2395% | 0.8602% | 1.481%
| Z-Score (Absolute) |  | 0.81 | 0.13 | 1.12 | 1.74 | 6.88 | 1.68 | 1.36 | 0.32 | 0.72 | 0.08 | NA | 0.51 | 0.77 | 0.64 | 0 | 0.43 | 0.71 | 1.28 | 3.18 | 0.24 | 0 | 0.16 | 0.32
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 555b (Indicative Value 0.073 ppm) | [28 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [24 empty] | 1
| No | No | 18 | 19 | 25 | [24 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [24 empty] | 3
3
| 1 | 1 | < 1 | NA | 0.08 | [24 empty] | 1
| 1 | 2 | < 1 | NA | 0.06 | [24 empty] | 20
| 1 | 3 | < 1 | NA | 0.07 | [24 empty] | 16
| 1 | 4 | < 1 | NA | 0.07 | [24 empty] | 0.073333
| 1 | 5 | < 1 | NA | 0.08 | [24 empty] | 28
| 1 | 6 | < 1 | NA | 0.08
| Mean |  |  |  | 0.073333
| Median |  |  |  | 0.075
| Std Dev. |  |  |  | 0.008165
| Rel.Std.Dev. |  |  |  | 11.13%
| PDM3 |  |  |  | 6.661e-14%
| Z-Score (Absolute) |  | 0.67 | NA | 0.67
| Analytical results for Ho in OREAS 555b (Certified Value 0.9 ppm) | [28 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.7 | 0.94 | 1.1 | 1.2992 | 1 | 1.2164 | 0.92513 | 0.85 | 0.72 | 0.77 | 1 | [16 empty] | 1
| 1 | 2 | 0.7 | 0.93 | 1 | 1.318 | 1 | 1.0366 | 0.90674 | 0.79 | 0.73 | 0.91 | 1 | [16 empty] | 3
| 1 | 3 | 0.6 | 0.92 | 1.1 | 1.2891 | 1 | 1.0098 | 0.90708 | 0.82 | 0.71 | 0.79 | 1.01 | [16 empty] | 16
| 1 | 4 | 0.6 | 0.93 | 0.8 | 1.2651 | 1 | 0.96445 | 0.93582 | 0.87 | 0.69 | 0.85 | 0.99 | [16 empty] | 0.902
| 1 | 5 | 0.6 | 0.96 | 1.1 | 1.2673 | 1 | 1.9348 | 0.95228 | 0.83 | 0.73 | 0.99 | 1.02 | [16 empty] | 108
| 1 | 6 | 0.6 | 0.95 | 1.1 | 1.331 | 1 | 0.97114 | 0.92876 | 0.81 | 0.75 | 0.85 | 0.98
| Mean |  | 0.63333 | 0.93833 | 1.0333 | 1.295 | 1 | 1.1889 | 0.92597 | 0.82833 | 0.72167 | 0.86 | 1
| Median |  | 0.6 | 0.935 | 1.1 | 1.2942 | 1 | 1.0232 | 0.92695 | 0.825 | 0.725 | 0.85 | 1
| Std Dev. |  | 0.05164 | 0.01472 | 0.12111 | 0.0266 | 0 | 0.37688 | 0.017461 | 0.028577 | 0.020412 | 0.080747 | 0.014142
| Rel.Std.Dev. |  | 8.154% | 1.569% | 11.72% | 2.054% | 0% | 31.7% | 1.886% | 3.45% | 2.829% | 9.389% | 1.414%
| PDM3 |  | -29.79% | 4.029% | 14.56% | 43.57% | 10.87% | 31.8% | 2.658% | -8.167% | -19.99% | -4.656% | 10.87%
| Z-Score (Absolute) |  | NA | 0.05 | NA | 2.78 | NA | 1.97 | 0.05 | 0.8 | 1.61 | 0.55 | 0.52
| Analytical results for In in OREAS 555b (Certified Value 2.51 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 19 | 20 | 22 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 2.44 | 2.18 | 2.52 | 2.7 | 2.62 | 2.4767 | 2.5 | 2.18 | 2.49 | 2.68 | 2.56 | 2.4294 | 2.59 | 2.512 | 2.7 | 2.55 | 2.207 | 2.03 | 2.63 | 2.301 | [7 empty] | 1
| 1 | 2 | 2.84 | 2.17 | 2.51 | 2.7 | 2.67 | 2.4808 | 2.61 | 2.19 | 2.69 | 2.63 | 2.66 | 2.4954 | 2.55 | 2.353 | 2.6 | 2.54 | 2.148 | 2.36 | 2.61 | 2.348 | [7 empty] | 4
| 1 | 3 | 2.4 | 2.29 | 2.59 | 2.7 | 2.73 | 2.5185 | 2.59 | 2.33 | 2.5 | 2.75 | 2.55 | 2.4303 | 2.5 | 2.434 | 2.8 | 2.48 | 2.101 | 2.14 | 2.58 | 2.306 | [7 empty] | 16
| 1 | 4 | 2.44 | 2.26 | 2.62 | 2.7 | 2.68 | 2.4993 | 2.54 | 2.11 | 2.6 | 2.72 | 2.56 | 2.4992 | 2.66 | 2.595 | 2.7 | 2.36 | 2.093 | 2.18 | 2.63 | 2.429 | [7 empty] | 2.506
| 1 | 5 | 2.48 | 2.16 | 2.73 | 2.5 | 2.79 | 2.5009 | 2.62 | 2.35 | 2.58 | 2.51 | 2.63 | 2.5319 | 2.55 | 2.477 | 2.8 | 2.39 | 2.015 | 2.37 | 2.58 | 2.379 | [7 empty] | 109
| 1 | 6 | 2.43 | 2.25 | 2.59 | 2.6 | 2.72 | 2.4806 | 2.55 | 2.28 | 2.58 | 2.61 | 2.62 | 2.5028 | 2.54 | 2.444 | 2.8 | 2.56 | 2.169 | 2.29 | 2.59 | 2.339
| Mean |  | 2.505 | 2.2183 | 2.5933 | 2.65 | 2.7017 | 2.4928 | 2.5683 | 2.24 | 2.5733 | 2.65 | 2.5967 | 2.4815 | 2.565 | 2.4692 | 2.7333 | 2.48 | 2.1222 | 2.2283 | 2.6033 | 2.3503
| Median |  | 2.44 | 2.215 | 2.59 | 2.7 | 2.7 | 2.49 | 2.57 | 2.235 | 2.58 | 2.655 | 2.59 | 2.4973 | 2.55 | 2.4605 | 2.75 | 2.51 | 2.1245 | 2.235 | 2.6 | 2.3435
| Std Dev. |  | 0.1661 | 0.054924 | 0.079666 | 0.083666 | 0.058452 | 0.016269 | 0.046224 | 0.094657 | 0.073121 | 0.086487 | 0.045898 | 0.042023 | 0.054681 | 0.081357 | 0.08165 | 0.086487 | 0.067588 | 0.13467 | 0.023381 | 0.048049
| Rel.Std.Dev. |  | 6.631% | 2.476% | 3.072% | 3.157% | 2.164% | 0.6526% | 1.8% | 4.226% | 2.841% | 3.264% | 1.768% | 1.693% | 2.132% | 3.295% | 2.987% | 3.487% | 3.185% | 6.044% | 0.8981% | 2.044%
| PDM3 |  | -0.04084% | -11.48% | 3.484% | 5.745% | 7.807% | -0.5281% | 2.486% | -10.62% | 2.686% | 5.745% | 3.617% | -0.9789% | 2.353% | -1.471% | 9.071% | -1.038% | -15.32% | -11.08% | 3.883% | -6.213%
| Z-Score (Absolute) |  | 0.3 | 3.18 | 0.59 | 1.16 | 1.68 | 0.42 | 0.34 | 2.97 | 0.39 | 1.16 | 0.62 | 0.54 | 0.3 | 0.66 | 1.99 | 0.55 | 4.15 | 3.08 | 0.69 | 1.86
| Analytical results for K in OREAS 555b (Certified Value 2.54 wt.%) | [28 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.48 | 2.23 | 2.52 | 2.5854 | 2.7 | 2.34 | 2.52 | 2.4 | 2.46 | 2.47 | 2.6 | 2.6108 | 2.175 | 2.4712 | 2.5539 | 2.54 | 2.4987 | 2.4583 | 2.51 | 2.38 | 2.51 | 2.56 | 2.39 | 2.6038 | 2.6 | 2.58 |  | 1
| 1 | 2 | 2.45 | 2.54 | 2.55 | 2.5518 | 2.76 | 2.28 | 2.52 | 2.44 | 2.57 | 2.54 | 2.57 | 2.7102 | 2.167 | 2.5144 | 2.5677 | 2.58 | 2.5207 | 2.4996 | 2.49 | 2.47 | 2.53 | 2.57 | 2.32 | 2.5878 | 2.58 | 2.59 |  | #N/A
| 1 | 3 | 2.42 | 1.64 | 2.56 | 2.5526 | 2.82 | 2.32 | 2.56 | 2.48 | 2.61 | 2.57 | 2.59 | 2.6864 | 2.208 | 2.6839 | 2.5088 | 2.61 | 2.4745 | 2.515 | 2.46 | 2.38 | 2.52 | 2.57 | 2.55 | 2.6082 | 2.59 | 2.52 |  | 16
| 1 | 4 | 2.48 | 2.31 | 2.55 | 2.5149 | 2.77 | 2.3 | 2.52 | 2.45 | 2.58 | 2.54 | 2.58 | 2.6759 | 2.217 | 2.4687 | 2.5317 | 2.68 | 2.5389 | 2.4734 | 2.48 | 2.45 | 2.5 | 2.52 | 2.6 | 2.574 | 2.58 | 2.6 |  | 2.5371
| 1 | 5 | 2.39 | 2.04 | 2.52 | 2.6395 | 2.75 | 2.3 | 2.49 | 2.48 | 2.56 | 2.52 | 2.6 | 2.6976 | 2.183 | 2.5736 | 2.5898 | 2.65 | 2.526 | 2.4708 | 2.42 | 2.38 | 2.5 | 2.63 | 2.55 | 2.6055 | 2.61 | 2.58 |  | 110
| 1 | 6 | 2.43 | 2.2 | 2.53 | 2.5286 | 2.74 | 2.29 | 2.46 | 2.42 | 2.52 | 2.59 | 2.59 | 2.6943 | 2.2 | 2.6791 | 2.5442 | 2.62 | 2.5112 | 2.4849 | 2.41 | 2.46 | 2.52 | 2.49 | 2.75 | 2.5964 | 2.59 | 2.56
| Mean |  | 2.4417 | 2.16 | 2.5383 | 2.5621 | 2.7567 | 2.305 | 2.5117 | 2.445 | 2.55 | 2.5383 | 2.5883 | 2.6792 | 2.1917 | 2.5651 | 2.5494 | 2.6133 | 2.5117 | 2.4837 | 2.4617 | 2.42 | 2.5133 | 2.5567 | 2.5267 | 2.5959 | 2.5917 | 2.5717
| Median |  | 2.44 | 2.215 | 2.54 | 2.5522 | 2.755 | 2.3 | 2.52 | 2.445 | 2.565 | 2.54 | 2.59 | 2.6904 | 2.1915 | 2.544 | 2.5491 | 2.615 | 2.5159 | 2.4791 | 2.47 | 2.415 | 2.515 | 2.565 | 2.55 | 2.6001 | 2.59 | 2.58
| Std Dev. |  | 0.035449 | 0.30272 | 0.017224 | 0.0449 | 0.039328 | 0.021679 | 0.033714 | 0.032094 | 0.052915 | 0.041673 | 0.01169 | 0.035409 | 0.019694 | 0.097834 | 0.028199 | 0.049666 | 0.022692 | 0.020744 | 0.039707 | 0.044272 | 0.012111 | 0.048028 | 0.15345 | 0.013048 | 0.01169 | 0.028577
| Rel.Std.Dev. |  | 1.452% | 14.01% | 0.6786% | 1.752% | 1.427% | 0.9405% | 1.342% | 1.313% | 2.075% | 1.642% | 0.4517% | 1.322% | 0.8986% | 3.814% | 1.106% | 1.9% | 0.9035% | 0.8352% | 1.613% | 1.829% | 0.4819% | 1.879% | 6.073% | 0.5026% | 0.4511% | 1.111%
| PDM3 |  | -3.76% | -14.86% | 0.04999% | 0.9881% | 8.656% | -9.147% | -1.001% | -3.629% | 0.5098% | 0.04999% | 2.021% | 5.602% | -13.61% | 1.107% | 0.4842% | 3.006% | -1.002% | -2.105% | -2.972% | -4.614% | -0.9354% | 0.7726% | -0.4099% | 2.321% | 2.152% | 1.364%
| Z-Score (Absolute) |  | 1.26 | 4.94 | 0 | 0.31 | 2.85 | 3.05 | 0.35 | 1.22 | 0.15 | 0 | 0.65 | 1.84 | 4.53 | 0.35 | 0.14 | 0.98 | 0.35 | 0.71 | 1 | 1.54 | 0.33 | 0.24 | 0.15 | 0.75 | 0.7 | 0.44
| Analytical results for La in OREAS 555b (Certified Value 72 ppm) | [28 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 0
1
| 1 | 1 | 66.28 | 73.6 | 53.6 | 74.19 | 76.6 | 70.4 | 75.581 | 67.9 | 74.8 | 68 | 71 | 79.6 | 54.92 | 72 | 67.6882 | 81.2 | 74.4 | 63.4 | 66 | 61.9 | 76 | 74.2 | 69.5 |  |  |  |  | 1
| 1 | 2 | 65.82 | 75.6 | 53.5 | 74.86 | 75.3 | 70.9 | 74.8472 | 72.1 | 65 | 70.9 | 80.6 | 77.7 | 53.29 | 74 | 67.9795 | 82.3 | 76.3 | 62.4 | 60 | 70.9 | 77 | 71.1 | 71.2 |  |  |  |  | 5
| 1 | 3 | 65.36 | 73.3 | 53.1 | 75 | 73.8 | 72.3 | 75.9539 | 71.3 | 68.1 | 65.2 | 72 | 79.3 | 54.28 | 73 | 67.1417 | 83.4 | 74.1 | 62.5 | 66 | 61.6 | 77 | 75.4 | 67.5 |  |  |  |  | 16
| 1 | 4 | 66.6 | 76.6 | 52 | 73.64 | 55.8 | 71.9 | 75.7843 | 71.7 | 63.7 | 72.2 | 71.3 | 74.3 | 54.37 | 72 | 69.3042 | 81.5 | 72.9 | 63.6 | 63 | 68 | 77 | 111 | 73.3 |  |  |  |  | 71.7949
| 1 | 5 | 64.87 | 76 | 53 | 76.09 | 73.8 | 73.6 | 76.1697 | 72.9 | 69.1 | 67.7 | 66.4 | 77.8 | 54.46 | 72 | 68.2534 | 81.3 | 73.9 | 62.1 | 65 | 75.4 | 80 | 74.1 | 69.2 |  |  |  |  | 111
| 1 | 6 | 65.23 | 73.6 | 52 | 78.1 | 74.9 | 73.9 | 76.3437 | 72.1 | 65.4 | 65.5 | 71.8 | 76.8 | 54.73 | 71 | 67.9743 | 78.5 | 75.8 | 63 | 67 | 69.7 | 78 | 74.2 | 73.3
| Mean |  | 65.6933 | 74.7833 | 52.8667 | 75.3133 | 71.7 | 72.1667 | 75.78 | 71.3333 | 67.6833 | 68.25 | 72.1833 | 77.5833 | 54.3417 | 72.3333 | 68.0569 | 81.3667 | 74.5667 | 62.8333 | 64.5 | 67.9167 | 77.5 | 80 | 70.6667
| Median |  | 65.59 | 74.6 | 53.05 | 74.93 | 74.35 | 72.1 | 75.8691 | 71.9 | 66.75 | 67.85 | 71.55 | 77.75 | 54.415 | 72 | 67.9769 | 81.4 | 74.25 | 62.75 | 65.5 | 68.85 | 77 | 74.2 | 70.35
| Std Dev. |  | 0.6612 | 1.4456 | 0.70899 | 1.5959 | 7.8593 | 1.4052 | 0.53096 | 1.7637 | 4.0276 | 2.8233 | 4.6184 | 1.9219 | 0.56743 | 1.0328 | 0.71845 | 1.6293 | 1.2644 | 0.59554 | 2.5884 | 5.3701 | 1.3784 | 15.2542 | 2.3534
| Rel.Std.Dev. |  | 1.006% | 1.933% | 1.341% | 2.119% | 10.96% | 1.947% | 0.7007% | 2.472% | 5.951% | 4.137% | 6.398% | 2.477% | 1.044% | 1.428% | 1.056% | 2.002% | 1.696% | 0.9478% | 4.013% | 7.907% | 1.779% | 19.07% | 3.33%
| PDM3 |  | -8.499% | 4.162% | -26.36% | 4.901% | -0.1322% | 0.5178% | 5.551% | -0.6429% | -5.727% | -4.938% | 0.541% | 8.062% | -24.31% | 0.7499% | -5.207% | 13.33% | 3.861% | -12.48% | -10.16% | -5.402% | 7.946% | 11.43% | -1.572%
| Z-Score (Absolute) |  | 1.07 | 0.55 | 3.36 | 0.64 | 0 | 0.08 | 0.73 | 0.07 | 0.72 | 0.61 | 0.09 | 1.05 | 3.09 | 0.11 | 0.65 | 1.72 | 0.51 | 1.58 | 1.28 | 0.67 | 1.03 | 1.48 | 0.18
| Analytical results for Li in OREAS 555b (Certified Value 59 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 61.08 | 65.7 | 59.5 | 63.2 | 58 | 60.1 | 52.7801 | 58.5 | 56.5 | 56.4 | 59.2 | 62 | 63.04 | 46 | 60.0613 | 60.7853 | 60.8 | 58.9008 | 55 | 56.7 | 59 | 56.4 | 59 | 49.9 | 59 | 57.6 | 50.7 | 1
| 1 | 2 | 60.29 | 59.7 | 60.5 | 61.7 | 57 | 62.7 | 53.9694 | 58 | 57.4 | 58.8 | 60.1 | 61.3 | 62.12 | 47 | 59.4303 | 60.2669 | 59.9 | 59.7399 | 56 | 58.3 | 61 | 53.7 | 57 | 47.4 | 60 | 57 | 52.1 | 33
| 1 | 3 | 59.23 | 63.4 | 60.9 | 62.5 | 56 | 62.6 | 53.8835 | 59.3 | 58.4 | 59.7 | 60.5 | 61.8 | 62.86 | 48 | 59.9656 | 60.8459 | 62.9 | 58.9232 | 56 | 59.4 | 58 | 58.2 | 57 | 53 | 59 | 58.3 | 50 | 16
| 1 | 4 | 61.88 | 66.4 | 59.7 | 61.7 | 55 | 63.2 | 53.1183 | 58.4 | 58.1 | 59.2 | 60.6 | 61.4 | 62.74 | 48 | 60.4717 | 60.7929 | 62.5 | 59.1563 | 55 | 59.6 | 60 | 57.8 | 59 | 55.7 | 59 | 56.4 | 54 | 59.3934
| 1 | 5 | 59.41 | 66.7 | 59.5 | 61.8 | 55 | 64.5 | 54.7718 | 58.2 | 58.6 | 58.4 | 59.8 | 61.9 | 62.88 | 47 | 61.2764 | 60.1069 | 61.4 | 59.7078 | 55 | 59.8 | 58 | 59.1 | 58 | 50.5 | 61 | 56.9 | 51.8 | 112
| 1 | 6 | 60.29 | 59.4 | 59.4 | 61.2 | 55 | 65.6 | 51.9772 | 57.7 | 57.2 | 57.4 | 61.7 | 61.4 | 62.93 | 47 | 61.2906 | 60.6834 | 61 | 60.0308 | 58 | 59.2 | 61 | 55.8 | 59 | 53 | 60 | 55.2 | 52
| Mean |  | 60.3633 | 63.55 | 59.9167 | 62.0167 | 56 | 63.1167 | 53.4167 | 58.35 | 57.7 | 58.3167 | 60.3167 | 61.6333 | 62.7617 | 47.1667 | 60.416 | 60.5802 | 61.4167 | 59.4098 | 55.8333 | 58.8333 | 59.5 | 56.8333 | 58.1667 | 51.5833 | 59.6667 | 56.9 | 51.7667
| Median |  | 60.29 | 64.55 | 59.6 | 61.75 | 55.5 | 62.95 | 53.5009 | 58.3 | 57.75 | 58.6 | 60.3 | 61.6 | 62.87 | 47 | 60.2665 | 60.7344 | 61.2 | 59.432 | 55.5 | 59.3 | 59.5 | 57.1 | 58.5 | 51.75 | 59.5 | 56.95 | 51.9
| Std Dev. |  | 1.0018 | 3.3086 | 0.62743 | 0.71391 | 1.2649 | 1.8777 | 0.99259 | 0.54681 | 0.80498 | 1.2205 | 0.84715 | 0.30111 | 0.32915 | 0.75277 | 0.74948 | 0.31327 | 1.1161 | 0.47823 | 1.169 | 1.1673 | 1.3784 | 1.95 | 0.98319 | 2.9144 | 0.8165 | 1.0583 | 1.3721
| Rel.Std.Dev. |  | 1.66% | 5.206% | 1.047% | 1.151% | 2.259% | 2.975% | 1.858% | 0.9371% | 1.395% | 2.093% | 1.405% | 0.4885% | 0.5244% | 1.596% | 1.241% | 0.5171% | 1.817% | 0.805% | 2.094% | 1.984% | 2.317% | 3.431% | 1.69% | 5.65% | 1.368% | 1.86% | 2.651%
| PDM3 |  | 1.633% | 6.998% | 0.8811% | 4.417% | -5.713% | 6.269% | -10.06% | -1.757% | -2.851% | -1.813% | 1.555% | 3.771% | 5.671% | -20.59% | 1.722% | 1.998% | 3.407% | 0.02767% | -5.994% | -0.9429% | 0.1795% | -4.31% | -2.065% | -13.15% | 0.4601% | -4.198% | -12.84%
| Z-Score (Absolute) |  | 0.32 | 1.39 | 0.17 | 0.88 | 1.15 | 1.25 | 2.01 | 0.36 | 0.57 | 0.37 | 0.3 | 0.75 | 1.13 | 4.11 | 0.34 | 0.39 | 0.67 | 0 | 1.2 | 0.19 | 0.03 | 0.87 | 0.42 | 2.63 | 0.09 | 0.84 | 2.57
| Analytical results for Lu in OREAS 555b (Certified Value 0.44 ppm) | [28 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 18 | 19 | 23 | 25 | 26 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [18 empty] | 2
3
| 1 | 1 | 0.4 | 0.44 | 0.5 | 0.47302 | 0.42932 | NA | 0.36 | 0.38 | 0.44 | [18 empty] | 1
| 1 | 2 | 0.4 | 0.46 | 0.4 | 0.48595 | 0.43708 | NA | 0.35 | 0.43 | 0.44 | [18 empty] | 25
| 1 | 3 | 0.4 | 0.44 | 0.4 | 0.46697 | 0.44039 | NA | 0.36 | 0.37 | 0.44 | [18 empty] | 16
| 1 | 4 | 0.4 | 0.46 | 0.3 | 0.4603 | 0.43661 | NA | 0.34 | 0.4 | 0.43 | [18 empty] | 0.44151
| 1 | 5 | 0.4 | 0.44 | 0.5 | 0.46782 | 0.45252 | NA | 0.34 | 0.44 | 0.43 | [18 empty] | 113
| 1 | 6 | 0.4 | 0.46 | 0.5 | 0.48328 | 0.46209 | NA | 0.35 | 0.41 | 0.44
| Mean |  | 0.4 | 0.45 | 0.43333 | 0.47289 | 0.443 |  | 0.35 | 0.405 | 0.43667
| Median |  | 0.4 | 0.45 | 0.45 | 0.47042 | 0.43874 |  | 0.35 | 0.405 | 0.44
| Std Dev. |  | 6.0809e-17 | 0.010954 | 0.08165 | 0.0099779 | 0.012038 |  | 0.0089443 | 0.027386 | 0.005164
| Rel.Std.Dev. |  | 1.52e-14% | 2.434% | 18.84% | 2.11% | 2.717% |  | 2.556% | 6.762% | 1.183%
| PDM3 |  | -9.402% | 1.923% | -1.852% | 7.107% | 0.3377% |  | -20.73% | -8.27% | -1.097%
| Z-Score (Absolute) |  | NA | 0.32 | NA | 1.03 | 0.1 | NA | 2.8 | 1.09 | 0.1
Results from laboratories 2 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 555b (Certified Value 3.07 wt.%) | [28 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.14 | 3.19 | 2.98 | 3.0465 | 3.01 | 3.03 | 3 | 2.95 | 2.9 | 3.01 | 3.17 | 3.1014 | 2.702 | 3.0866 | 3.2406 | 3 | 3.0331 | 3.0537 | 3.06 | 2.89 | 2.73 | 3.128 | 2.86 | 3.0662 | 3.07 | 3.16 |  | 1
| 1 | 2 | 3.11 | 3.15 | 3.03 | 2.9946 | 3.07 | 2.98 | 2.97 | 2.99 | 3.02 | 3.07 | 3.13 | 3.2119 | 2.678 | 3.2403 | 3.276 | 3.1 | 3.0577 | 3.0836 | 3.03 | 2.93 | 2.83 | 3.197 | 2.72 | 3.0888 | 3.04 | 3.17 |  | #N/A
| 1 | 3 | 3.08 | 3.14 | 3.01 | 3.0006 | 3.12 | 3.06 | 3.06 | 3.03 | 3.04 | 3.1 | 3.15 | 3.1494 | 2.744 | 3.2811 | 3.217 | 3.08 | 2.9943 | 3.1171 | 3 | 2.93 | 2.78 | 3.176 | 2.97 | 3.0484 | 3.06 | 3.1 |  | 16
| 1 | 4 | 3.15 | 3.26 | 3.06 | 2.9321 | 3.07 | 2.99 | 3.01 | 3.01 | 3.03 | 3.06 | 3.13 | 3.1386 | 2.726 | 3.0777 | 3.2098 | 3.15 | 3.0436 | 3.0629 | 3.03 | 2.79 | 2.79 | 3.093 | 3.08 | 3.0441 | 3.02 | 3.19 |  | 3.0723
| 1 | 5 | 3.03 | 3.28 | 2.99 | 3.0852 | 3.06 | 3.08 | 2.98 | 3.04 | 3.02 | 3.03 | 3.18 | 3.1455 | 2.708 | 3.1808 | 3.2584 | 3.1 | 3.0783 | 3.0611 | 2.97 | 2.88 | 2.83 | 3.29 | 3 | 3.1249 | 3.08 | 3.16 |  | 114
| 1 | 6 | 3.05 | 3.14 | 2.96 | 2.9715 | 3.06 | 2.97 | 2.96 | 2.98 | 2.98 | 3.12 | 3.15 | 3.1674 | 2.684 | 3.0637 | 3.2245 | 3.15 | 3.0557 | 3.0581 | 2.94 | 2.86 | 2.92 | 3.117 | 3.25 | 3.073 | 3.06 | 3.14
| Mean |  | 3.0933 | 3.1933 | 3.005 | 3.0051 | 3.065 | 3.0183 | 2.9967 | 3 | 2.9983 | 3.065 | 3.1517 | 3.1524 | 2.707 | 3.1551 | 3.2377 | 3.0967 | 3.0438 | 3.0727 | 3.005 | 2.88 | 2.8133 | 3.1668 | 2.98 | 3.0742 | 3.055 | 3.1533
| Median |  | 3.095 | 3.17 | 3 | 2.9976 | 3.065 | 3.01 | 2.99 | 3 | 3.02 | 3.065 | 3.15 | 3.1475 | 2.705 | 3.1337 | 3.2325 | 3.1 | 3.0496 | 3.062 | 3.015 | 2.885 | 2.81 | 3.152 | 2.985 | 3.0696 | 3.06 | 3.16
| Std Dev. |  | 0.048442 | 0.062503 | 0.036194 | 0.054256 | 0.035071 | 0.04535 | 0.036148 | 0.033466 | 0.052313 | 0.041352 | 0.020412 | 0.036368 | 0.025004 | 0.092547 | 0.025632 | 0.055377 | 0.028619 | 0.024062 | 0.044159 | 0.052154 | 0.064083 | 0.071564 | 0.18188 | 0.029735 | 0.021679 | 0.030768
| Rel.Std.Dev. |  | 1.566% | 1.957% | 1.204% | 1.805% | 1.144% | 1.503% | 1.206% | 1.116% | 1.745% | 1.349% | 0.6477% | 1.154% | 0.9237% | 2.933% | 0.7917% | 1.788% | 0.9402% | 0.7831% | 1.47% | 1.811% | 2.278% | 2.26% | 6.103% | 0.9672% | 0.7096% | 0.9757%
| PDM3 |  | 0.6833% | 3.938% | -2.192% | -2.189% | -0.2389% | -1.758% | -2.463% | -2.355% | -2.409% | -0.2389% | 2.582% | 2.605% | -11.89% | 2.692% | 5.383% | 0.7918% | -0.9293% | 0.01335% | -2.192% | -6.26% | -8.43% | 3.076% | -3.006% | 0.06163% | -0.5644% | 2.636%
| Z-Score (Absolute) |  | 0.37 | 1.48 | 0.61 | 0.61 | 0.06 | 0.46 | 0.7 | 0.67 | 0.68 | 0.06 | 1.02 | 1.02 | 3.92 | 1.05 | 1.97 | 0.41 | 0.18 | 0.14 | 0.61 | 2 | 2.73 | 1.18 | 0.89 | 0.16 | 0.06 | 1.03
| Analytical results for Mn in OREAS 555b (Certified Value 0.089 wt.%) | [28 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.08675 | 0.1025 | 0.0871 | 0.092 | 0.0904 | 0.0873 | 0.0846 | 0.085 | 0.0858 | 0.0846 | 0.093 | 0.0858 | 0.077 | 0.09565 | 0.083795 | 0.0859 | 0.089446 | 0.0937 | 0.09 | 0.0872 | 0.0961 | 0.0851 | 0.075 | 0.092 | 0.0932 | 0.0908 |  | 1
| 1 | 2 | 0.08436 | 0.0899 | 0.0875 | 0.0907 | 0.092 | 0.086 | 0.0849 | 0.0908 | 0.0889 | 0.0846 | 0.0908 | 0.0917 | 0.077 | 0.094394 | 0.083966 | 0.0866 | 0.089223 | 0.0913 | 0.088 | 0.0891 | 0.0942 | 0.0894 | 0.0811 | 0.0912 | 0.0909 | 0.0883 |  | #N/A
| 1 | 3 | 0.07999 | 0.0948 | 0.0871 | 0.0884 | 0.0902 | 0.0888 | 0.0859 | 0.0881 | 0.0912 | 0.0891 | 0.0919 | 0.0879 | 0.085 | 0.095052 | 0.083329 | 0.0877 | 0.0896 | 0.0917 | 0.0902 | 0.0864 | 0.0944 | 0.0888 | 0.0845 | 0.0901 | 0.0938 | 0.0869 |  | 16
| 1 | 4 | 0.08251 | 0.0983 | 0.0876 | 0.0863 | 0.0704 | 0.0869 | 0.0849 | 0.0897 | 0.0903 | 0.087 | 0.0875 | 0.0884 | 0.077 | 0.093331 | 0.083058 | 0.0843 | 0.089266 | 0.091 | 0.0874 | 0.0888 | 0.0943 | 0.0887 | 0.0856 | 0.0909 | 0.0882 | 0.0885 |  | 0.089023
| 1 | 5 | 0.08066 | 0.0965 | 0.0871 | 0.095 | 0.0907 | 0.0889 | 0.0863 | 0.086 | 0.0963 | 0.0863 | 0.0902 | 0.0885 | 0.077 | 0.091642 | 0.083712 | 0.0892 | 0.090235 | 0.093 | 0.09 | 0.0891 | 0.0937 | 0.0903 | 0.0792 | 0.09 | 0.0961 | 0.0857 |  | 115
| 1 | 6 | 0.08047 | 0.0933 | 0.0871 | 0.0877 | 0.092 | 0.0887 | 0.0861 | 0.0914 | 0.0858 | 0.0871 | 0.0911 | 0.0886 | 0.077 | 0.092291 | 0.08301 | 0.09 | 0.089643 | 0.0921 | 0.0889 | 0.0895 | 0.0952 | 0.0894 | 0.0855 | 0.0931 | 0.0943 | 0.0901
| Mean |  | 0.082457 | 0.095883 | 0.08725 | 0.090017 | 0.087617 | 0.087767 | 0.08545 | 0.0885 | 0.089717 | 0.08645 | 0.09075 | 0.088483 | 0.078333 | 0.093727 | 0.083478 | 0.087283 | 0.089569 | 0.092133 | 0.089083 | 0.08835 | 0.09465 | 0.088617 | 0.081817 | 0.091217 | 0.09275 | 0.088383
| Median |  | 0.081585 | 0.09565 | 0.0871 | 0.08955 | 0.09055 | 0.088 | 0.0854 | 0.0889 | 0.0896 | 0.08665 | 0.09095 | 0.08845 | 0.077 | 0.093862 | 0.08352 | 0.08715 | 0.089523 | 0.0919 | 0.08945 | 0.08895 | 0.09435 | 0.0891 | 0.0828 | 0.09105 | 0.0935 | 0.0884
| Std Dev. |  | 0.0026577 | 0.0043278 | 0.00023452 | 0.0031946 | 0.0084708 | 0.0012094 | 0.00073144 | 0.0026 | 0.0039301 | 0.0017097 | 0.0018641 | 0.0018925 | 0.003266 | 0.001579 | 0.00040274 | 0.0021217 | 0.00036799 | 0.0010367 | 0.0011805 | 0.001247 | 0.00085965 | 0.0018148 | 0.0042168 | 0.0011822 | 0.0027948 | 0.0019083
| Rel.Std.Dev. |  | 3.223% | 4.514% | 0.2688% | 3.549% | 9.668% | 1.378% | 0.856% | 2.938% | 4.381% | 1.978% | 2.054% | 2.139% | 4.169% | 1.685% | 0.4825% | 2.431% | 0.4108% | 1.125% | 1.325% | 1.411% | 0.9082% | 2.048% | 5.154% | 1.296% | 3.013% | 2.159%
| PDM3 |  | -7.376% | 7.706% | -1.992% | 1.116% | -1.58% | -1.412% | -4.014% | -0.5879% | 0.7788% | -2.891% | 1.94% | -0.6066% | -12.01% | 5.283% | -6.229% | -1.955% | 0.6128% | 3.493% | 0.06738% | -0.7564% | 6.32% | -0.4568% | -8.095% | 2.464% | 4.186% | -0.7189%
| Z-Score (Absolute) |  | 2.28 | 2.79 | 0.47 | 0.58 | 0.33 | 0.27 | 1.15 | 0 | 0.46 | 0.77 | 0.85 | 0 | 3.69 | 1.98 | 1.9 | 0.46 | 0.41 | 1.38 | 0.22 | 0.05 | 2.33 | 0.05 | 2.52 | 1.03 | 1.61 | 0.04
| Analytical results for Mo in OREAS 555b (Certified Value 17.1 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 4.82 | 14.89 | 15.9 | 17.5 | 18 | 16.8 | 17.6358 | 15.75 | 15.85 | 16.9 | 18.25 | 18.8 | 15.44 | 15 | 16.6124 | 17.2 | 16.5651 | 17.4 | 17.23 | 15 | 18.9 | 13.35 | 14.5 | 18.5 | 16.8 | 16.78 |  | 1
| 1 | 2 | 3.45 | 15.32 | 16.3 | 17 | 18 | 17.6 | 17.3406 | 16.6 | 16.15 | 17.8 | 18 | 18.45 | 16.76 | 8 | 16.6328 | 17.1 | 16.8109 | 16.4 | 16.65 | 17 | 18.8 | 13.2 | 17.1 | 18.2 | 16.5 | 16.7 |  | 7
| 1 | 3 | 4.33 | 14.91 | 16.3 | 17.2 | 17 | 17.4 | 17.3867 | 16.55 | 16.2 | 16.65 | 18.45 | 18.9 | 16.82 | 18 | 16.7294 | 17.5 | 16.4404 | 17.1 | 17.43 | 16 | 18 | 13.3 | 15 | 18 | 16.8 | 17.02 |  | 16
| 1 | 4 | 4.68 | 14.39 | 17.2 | 17.5 | 18 | 17.8 | 17.6866 | 16.15 | 15.7 | 17.9 | 18.25 | 17 | 16.37 | 10 | 17.1948 | 17.8 | 16.7413 | 18.2 | 17.13 | 15 | 17.9 | 13.26 | 15.9 | 18 | 16.8 | 17.65 |  | 17.0696
| 1 | 5 | 3.81 | 15.44 | 17.1 | 17.7 | 17 | 18 | 17.3532 | 16.6 | 17.1 | 17.05 | 17.05 | 18.2 | 16.48 | 12 | 17.3194 | 17.8 | 16.7572 | 17.5 | 17.25 | 16 | 18.2 | 13.32 | 17.5 | 18.2 | 16.7 | 17.02 |  | 116
| 1 | 6 | 4.51 | 14.64 | 16.7 | 18.1 | 17 | 18.2 | 17.1625 | 16.5 | 16.1 | 17.25 | 18.2 | 18.3 | 16.09 | 16 | 17.0894 | 17 | 16.6413 | 17.5 | 17.35 | 17 | 18.5 | 13.61 | 15.7 | 17.9 | 16.6 | 16.85
| Mean |  | 4.2667 | 14.9317 | 16.5833 | 17.5 | 17.5 | 17.6333 | 17.4276 | 16.3583 | 16.1833 | 17.2583 | 18.0333 | 18.275 | 16.3267 | 13.1667 | 16.9297 | 17.4 | 16.6594 | 17.35 | 17.1733 | 16 | 18.3833 | 13.34 | 15.95 | 18.1333 | 16.7 | 17.0033
| Median |  | 4.42 | 14.9 | 16.5 | 17.5 | 17.5 | 17.7 | 17.37 | 16.525 | 16.125 | 17.15 | 18.225 | 18.375 | 16.425 | 13.5 | 16.9094 | 17.35 | 16.6913 | 17.45 | 17.24 | 16 | 18.35 | 13.31 | 15.8 | 18.1 | 16.75 | 16.935
| Std Dev. |  | 0.53212 | 0.39746 | 0.50761 | 0.38471 | 0.54772 | 0.49666 | 0.19771 | 0.34266 | 0.48854 | 0.49942 | 0.50266 | 0.68246 | 0.50965 | 3.8166 | 0.30873 | 0.35214 | 0.13885 | 0.58907 | 0.27638 | 0.89443 | 0.41673 | 0.14213 | 1.1658 | 0.21602 | 0.12649 | 0.3418
| Rel.Std.Dev. |  | 12.47% | 2.662% | 3.061% | 2.198% | 3.13% | 2.817% | 1.134% | 2.095% | 3.019% | 2.894% | 2.787% | 3.734% | 3.122% | 28.99% | 1.824% | 2.024% | 0.8335% | 3.395% | 1.609% | 5.59% | 2.267% | 1.065% | 7.309% | 1.191% | 0.7574% | 2.01%
| PDM3 |  | -75% | -12.52% | -2.849% | 2.521% | 2.521% | 3.303% | 2.097% | -4.167% | -5.192% | 1.106% | 5.646% | 7.062% | -4.352% | -22.86% | -0.8197% | 1.936% | -2.403% | 1.643% | 0.6077% | -6.266% | 7.696% | -21.85% | -6.559% | 6.232% | -2.165% | -0.3882%
| Z-Score (Absolute) |  | 13.63 | 2.22 | 0.45 | 0.53 | NA | 0.67 | 0.45 | 0.69 | 0.88 | 0.27 | 1.1 | 1.36 | 0.72 | NA | 0.08 | 0.42 | 0.37 | 0.37 | 0.18 | NA | 1.48 | 3.92 | 1.13 | 1.21 | 0.32 | 0
Results from laboratories 5, 14 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 555b (Certified Value 0.052 wt.%) | [28 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.06 | 0.065 | 0.04 | 0.0481 | 0.05 | 0.05 | 0.06 | 0.04 | 0.05 | 0.05 | 0.06 | 0.0541 | 0.045 | 0.062882 | 0.075213 | 0.046 | 0.050348 | NA | 0.08 | 0.46 | 0.12 | 0.0263 | 0.0475 | 0.05 | 0.05 |  |  | 1
| 1 | 2 | 0.05 | 0.053 | 0.04 | 0.0476 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.0557 | 0.045 | 0.061431 | 0.075058 | 0.047 | 0.05164 | NA | 0.08 | 0.51 | 0.11 | 0.0251 | 0.0477 | 0.05 | 0.05 |  |  | 17
| 1 | 3 | 0.06 | 0.064 | 0.04 | 0.0476 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.0551 | 0.045 | 0.0654 | 0.075164 | 0.048 | 0.051652 | NA | 0.08 | 0.43 | 0.11 | 0.0267 | 0.0478 | 0.05 | 0.05 |  |  | 16
| 1 | 4 | 0.06 | 0.067 | 0.04 | 0.046 | 0.04 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.0552 | 0.045 | 0.068748 | 0.075577 | 0.047 | 0.051842 | NA | 0.08 | 0.44 | 0.12 | 0.0311 | 0.0478 | 0.05 | 0.05 |  |  | 0.051973
| 1 | 5 | 0.05 | 0.067 | 0.04 | 0.049 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.0554 | 0.045 | 0.067718 | 0.075507 | 0.047 | 0.051411 | NA | 0.08 | 0.52 | 0.12 | 0.0248 | 0.0491 | 0.05 | 0.05 |  |  | 117
| 1 | 6 | 0.06 | 0.064 | 0.04 | 0.047 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.0554 | 0.045 | 0.062743 | 0.075038 | 0.049 | 0.051579 | NA | 0.08 | 0.51 | 0.11 | 0.027 | 0.048 | 0.05 | 0.05
| Mean |  | 0.056667 | 0.063333 | 0.04 | 0.04755 | 0.048333 | 0.05 | 0.06 | 0.048333 | 0.05 | 0.05 | 0.06 | 0.05515 | 0.045 | 0.06482 | 0.075259 | 0.047333 | 0.051412 |  | 0.08 | 0.47833 | 0.115 | 0.026833 | 0.047983 | 0.05 | 0.05
| Median |  | 0.06 | 0.0645 | 0.04 | 0.0476 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.0553 | 0.045 | 0.064141 | 0.075188 | 0.047 | 0.051609 |  | 0.08 | 0.485 | 0.115 | 0.0265 | 0.0478 | 0.05 | 0.05
| Std Dev. |  | 0.005164 | 0.0052409 | 0 | 0.0010114 | 0.0040825 | 7.6012e-18 | 0 | 0.0040825 | 7.6012e-18 | 7.6012e-18 | 0 | 0.00055408 | 7.6012e-18 | 0.002957 | 0.00022935 | 0.0010328 | 0.00053931 |  | 0 | 0.039707 | 0.0054772 | 0.002266 | 0.00057067 | 7.6012e-18 | 7.6012e-18
| Rel.Std.Dev. |  | 9.113% | 8.275% | 0% | 2.127% | 8.447% | 1.52e-14% | 0% | 8.447% | 1.52e-14% | 1.52e-14% | 0% | 1.005% | 1.689e-14% | 4.562% | 0.3047% | 2.182% | 1.049% |  | 0% | 8.301% | 4.763% | 8.445% | 1.189% | 1.52e-14% | 1.52e-14%
| PDM3 |  | 9.03% | 21.86% | -23.04% | -8.511% | -7.004% | -3.797% | 15.44% | -7.004% | -3.797% | -3.797% | 15.44% | 6.112% | -13.42% | 24.72% | 44.8% | -8.928% | -1.08% |  | 53.93% | 820.3% | 121.3% | -48.37% | -7.677% | -3.797% | -3.797%
| Z-Score (Absolute) |  | 0.85 | 1.69 | 1.27 | 0.31 | 0.21 | 0 | 1.27 | 0.21 | 0 | 0 | 1.27 | 0.65 | 0.7 | 1.88 | 3.2 | 0.34 | 0.18 | NA | 3.8 | 54.29 | 8.24 | 2.94 | 0.26 | 0 | 0
| Analytical results for Nb in OREAS 555b (Certified Value 17.1 ppm) | [28 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 16.38 | 6.8 | 14.73 | 25.8 | 9.8 | 25.4095 | 19.2 | 17.4 | 18 | 17.8 | 22.6 | 46 | 20.6274 | 9.2 | 17.4098 | NA | 19 | 7 | 15.9 | 13.5 | 16.57 | 13 | 18.9 |  |  |  |  | 1
| 1 | 2 | 17.03 | 6.74 | 15.26 | 25.6 | 10.3 | 25.5762 | 20 | 17.1 | 19.2 | 17.2 | 22.9 | 45 | 19.8717 | 9.6 | 17.7365 | NA | 20.1 | 6 | 15 | 14 | 17.16 | 12.6 | 19.4 |  |  |  |  | 9
| 1 | 3 | 16.4 | 6.51 | 14.69 | 25.4 | 10.3 | 25.5402 | 20.4 | 17.7 | 18.2 | 17.4 | 23.1 | 46 | 20.1668 | 9.5 | 17.1352 | NA | 19.6 | 7 | 16.6 | 12.6 | 16.6 | 12.8 | 18.5 |  |  |  |  | 16
| 1 | 4 | 16.71 | 6.5 | 13.07 | 18.4 | 10.1 | 25.7415 | 19.5 | 17 | 19.2 | 17.3 | 21.3 | 47 | 20.7583 | 9.7 | 17.7886 | NA | 19.3 | 6 | 15.2 | 15 | 17.44 | 12.4 | 20 |  |  |  |  | 17.0885
| 1 | 5 | 16.96 | 6.51 | 15.44 | 24.4 | 9.9 | 25.6736 | 20 | 17.2 | 18.4 | 15.7 | 22.1 | 43 | 19.7951 | 9.6 | 17.8779 | NA | 20 | 7 | 16.1 | 15.3 | 17.91 | 12.8 | 19.4 |  |  |  |  | 118
| 1 | 6 | 16.57 | 6.98 | 13.3 | 24.3 | 9.7 | 25.2126 | 20.1 | 17 | 18.6 | 17.4 | 22 | 44 | 19.4086 | 9.6 | 17.7903 | NA | 20 | 6 | 15.9 | 14.1 | 16.77 | 12.5 | 19.1
| Mean |  | 16.675 | 6.6733 | 14.415 | 23.9833 | 10.0167 | 25.5256 | 19.8667 | 17.2333 | 18.6 | 17.1333 | 22.3333 | 45.1667 | 20.1046 | 9.5333 | 17.623 |  | 19.6667 | 6.5 | 15.7833 | 14.0833 | 17.075 | 12.6833 | 19.2167
| Median |  | 16.64 | 6.625 | 14.71 | 24.9 | 10 | 25.5582 | 20 | 17.15 | 18.5 | 17.35 | 22.35 | 45.5 | 20.0193 | 9.6 | 17.7625 |  | 19.8 | 6.5 | 15.9 | 14.05 | 16.965 | 12.7 | 19.25
| Std Dev. |  | 0.27646 | 0.19896 | 0.99917 | 2.806 | 0.25626 | 0.19116 | 0.43665 | 0.27325 | 0.50596 | 0.73121 | 0.66533 | 1.472 | 0.51749 | 0.17512 | 0.28866 |  | 0.44572 | 0.54772 | 0.59133 | 0.98675 | 0.53129 | 0.22286 | 0.51153
| Rel.Std.Dev. |  | 1.658% | 2.981% | 6.931% | 11.7% | 2.558% | 0.7489% | 2.198% | 1.586% | 2.72% | 4.268% | 2.979% | 3.259% | 2.574% | 1.837% | 1.638% |  | 2.266% | 8.427% | 3.747% | 7.006% | 3.112% | 1.757% | 2.662%
| PDM3 |  | -2.42% | -60.95% | -15.65% | 40.35% | -41.38% | 49.37% | 16.26% | 0.8474% | 8.845% | 0.2622% | 30.69% | 164.3% | 17.65% | -44.21% | 3.128% |  | 15.09% | -61.96% | -7.638% | -17.59% | -0.07919% | -25.78% | 12.45%
| Z-Score (Absolute) |  | 0.13 | 2.6 | 0.68 | 1.68 | 1.77 | 2.06 | 0.66 | 0.01 | 0.35 | 0.01 | 1.27 | NA | 0.72 | 1.89 | 0.11 | NA | 0.61 | NA | 0.35 | 0.77 | 0.03 | 1.11 | 0.5
Results from laboratories 14 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 555b (Certified Value 64 ppm) | [28 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 0
1
| 1 | 1 | 43.4 | 62.99 | 66.2 | 62.8767 | 66 | 68.6786 | 62.2161 | 63.62 | 57.1 | 54.9 | 67.79 | [16 empty] | 1
| 1 | 2 | 43.8 | 62.92 | 66.1 | 63.2882 | 65 | 65.0391 | 61.7681 | 59.66 | 54.5 | 64.2 | 66.09 | [16 empty] | 10
| 1 | 3 | 45.9 | 62.15 | 64.3 | 63.1052 | 67 | 64.4669 | 60.5076 | 62.31 | 55.4 | 55.1 | 68.03 | [16 empty] | 16
| 1 | 4 | 52.4 | 63.65 | 62.5 | 62.4106 | 66 | 64.2891 | 63.1637 | 65.91 | 53.5 | 59.9 | 66.67 | [16 empty] | 64.0192
| 1 | 5 | 50.7 | 63.44 | 63.4 | 62.8387 | 65 | 64.5024 | 62.7326 | 63.2 | 55.4 | 66.6 | 67.55 | [16 empty] | 119
| 1 | 6 | 44.1 | 65.15 | 63.9 | 63.6847 | 66 | 63.5166 | 62.3268 | 61.6 | 57.4 | 61.9 | 66.91
| Mean |  | 46.7167 | 63.3833 | 64.4 | 63.034 | 65.8333 | 65.0821 | 62.1192 | 62.7167 | 55.55 | 60.4333 | 67.1733
| Median |  | 45 | 63.215 | 64.1 | 62.9909 | 66 | 64.4846 | 62.2715 | 62.755 | 55.4 | 60.9 | 67.23
| Std Dev. |  | 3.8778 | 1.0086 | 1.4832 | 0.43466 | 0.75277 | 1.8292 | 0.92089 | 2.0977 | 1.495 | 4.7689 | 0.74193
| Rel.Std.Dev. |  | 8.301% | 1.591% | 2.303% | 0.6896% | 1.143% | 2.811% | 1.482% | 3.345% | 2.691% | 7.891% | 1.105%
| PDM3 |  | -27.03% | -0.9932% | 0.5948% | -1.539% | 2.834% | 1.66% | -2.968% | -2.035% | -13.23% | -5.601% | 4.927%
| Z-Score (Absolute) |  | 5.37 | 0.12 | 0.45 | 0 | 0.92 | 0.67 | 0.3 | 0.1 | 2.46 | 0.86 | 1.36
| Analytical results for Ni in OREAS 555b (Certified Value 220 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 207.4 | 232.2 | 221 | 216.8 | 232 | 212 | 229.1114 | 218 | 211 | 212 | 217 | 226 | 194 | 174 | 227.9154 | 220.25 | 229 | 179.1112 | 214.2 | 246 | 201 | 209.4 | 152 | 188 | 224 | 235.4 | 230.9 | 1
| 1 | 2 | 200.7 | 217.4 | 224 | 210.7 | 222 | 211 | 223.1776 | 219 | 226 | 219 | 215 | 223 | 198 | 181 | 228.7521 | 229.16 | 233 | 176.3163 | 209.9 | 234 | 201 | 203.7 | 149 | 236 | 225 | 227.4 | 224.7 | 40
| 1 | 3 | 194.8 | 219.5 | 224 | 211.5 | 216 | 228 | 225.137 | 217 | 219 | 226 | 227 | 229 | 198 | 185 | 225.2992 | 222.44 | 230 | 179.719 | 210.7 | 242 | 203 | 202.5 | 147 | 202 | 220 | 232 | 223.1 | 16
| 1 | 4 | 194.3 | 225.8 | 226 | 216.9 | 213 | 214 | 237.8335 | 215 | 221 | 224 | 221 | 218 | 196 | 187 | 230.2099 | 220.65 | 235 | 178.199 | 208.6 | 237 | 213 | 198.6 | 141 | 207 | 221 | 226.6 | 214.7 | 220.1353
| 1 | 5 | 195 | 222.2 | 221 | 226 | 219 | 229 | 222.5753 | 221 | 214 | 238 | 220 | 225 | 196 | 172 | 223.9486 | 221.61 | 230 | 184.2438 | 212.2 | 241 | 194 | 200.3 | 140 | 234 | 221 | 241.1 | 220.6 | 120
| 1 | 6 | 196.6 | 212 | 221 | 215.8 | 214 | 221 | 226.7405 | 215 | 222 | 214 | 220 | 226 | 199 | 168 | 223.1267 | 222.3 | 230 | 180.3447 | 210.6 | 245 | 210 | 201.9 | 148 | 215 | 229 | 233.5 | 229.8
| Mean |  | 198.1333 | 221.5167 | 222.8333 | 216.2833 | 219.3333 | 219.1667 | 227.4292 | 217.5 | 218.8333 | 222.1667 | 220 | 224.5 | 196.8333 | 177.8333 | 226.542 | 222.735 | 231.1667 | 179.6557 | 211.0333 | 240.8333 | 203.6667 | 202.7333 | 146.1667 | 213.6667 | 223.3333 | 232.6667 | 223.9667
| Median |  | 195.8 | 220.85 | 222.5 | 216.3 | 217.5 | 217.5 | 225.9387 | 217.5 | 220 | 221.5 | 220 | 225.5 | 197 | 177.5 | 226.6073 | 221.955 | 230 | 179.4151 | 210.65 | 241.5 | 202 | 202.2 | 147.5 | 211 | 222.5 | 232.75 | 223.9
| Std Dev. |  | 5.1075 | 6.9933 | 2.137 | 5.4653 | 7.0333 | 8.0353 | 5.6293 | 2.3452 | 5.4924 | 9.4745 | 4.0988 | 3.7283 | 1.8348 | 7.6267 | 2.834 | 3.2663 | 2.3166 | 2.6495 | 1.9439 | 4.6224 | 6.8605 | 3.7184 | 4.7081 | 18.7261 | 3.3862 | 5.3717 | 6.0099
| Rel.Std.Dev. |  | 2.578% | 3.157% | 0.959% | 2.527% | 3.207% | 3.666% | 2.475% | 1.078% | 2.51% | 4.265% | 1.863% | 1.661% | 0.9322% | 4.289% | 1.251% | 1.466% | 1.002% | 1.475% | 0.9211% | 1.919% | 3.369% | 1.834% | 3.221% | 8.764% | 1.516% | 2.309% | 2.683%
| PDM3 |  | -9.995% | 0.6275% | 1.226% | -1.75% | -0.3643% | -0.44% | 3.313% | -1.197% | -0.5914% | 0.9228% | -0.06144% | 1.983% | -10.59% | -19.22% | 2.91% | 1.181% | 5.011% | -18.39% | -4.135% | 9.402% | -7.481% | -7.905% | -33.6% | -2.938% | 1.453% | 5.693% | 1.74%
| Z-Score (Absolute) |  | 2.52 | 0.26 | 0.42 | 0.36 | 0 | 0.02 | 0.96 | 0.22 | 0.06 | 0.34 | 0.08 | 0.61 | 2.68 | 4.94 | 0.86 | 0.4 | 1.41 | 4.72 | 0.99 | 2.56 | 1.86 | 1.98 | 8.71 | 0.67 | 0.48 | 1.59 | 0.55
| Analytical results for P in OREAS 555b (Certified Value 0.035 wt.%) | [28 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.03 | 0.038 | 0.0518 | 0.0344 | 0.041 | 0.0563 | 0.034 | 0.033 | 0.032 | 0.035 | 0.037 | 0.0696 | 0.127 | 0.034872 | 0.038 | 0.033586 | 0.031 | 0.03 | 0.03 | 0.034 | 0.0412 | 0.0349 | 0.03 | 0.0361 |  |  |  | 1
| 1 | 2 | 0.03 | 0.036 | 0.0521 | 0.0338 | 0.041 | 0.0558 | 0.034 | 0.033 | 0.035 | 0.035 | 0.036 | 0.0681 | 0.109 | 0.035396 | 0.036 | 0.033148 | 0.0315 | 0.03 | 0.03 | 0.034 | 0.0446 | 0.0352 | 0.03 | 0.0361 |  |  |  | #N/A
| 1 | 3 | 0.03 | 0.038 | 0.052 | 0.0343 | 0.041 | 0.0576 | 0.034 | 0.034 | 0.033 | 0.036 | 0.037 | 0.0695 | 0.1 | 0.036065 | 0.037 | 0.032501 | 0.0317 | 0.03 | 0.03 | 0.034 | 0.0414 | 0.0346 | 0.03 | 0.035 |  |  |  | 16
| 1 | 4 | 0.03 | 0.04 | 0.0519 | 0.0338 | 0.039 | 0.0584 | 0.034 | 0.034 | 0.034 | 0.035 | 0.037 | 0.0693 | 0.1 | 0.034782 | 0.039 | 0.033185 | 0.0316 | 0.03 | 0.03 | 0.033 | 0.0408 | 0.0348 | 0.03 | 0.0359 |  |  |  | 0.034577
| 1 | 5 | 0.03 | 0.04 | 0.0522 | 0.0346 | 0.04 | 0.0575 | 0.034 | 0.034 | 0.033 | 0.035 | 0.037 | 0.0694 | 0.118 | 0.036378 | 0.038 | 0.033466 | 0.0319 | 0.03 | 0.03 | 0.034 | 0.0406 | 0.0359 | 0.04 | 0.0356 |  |  |  | 121
| 1 | 6 | 0.04 | 0.039 | 0.0523 | 0.0343 | 0.04 | 0.0581 | 0.034 | 0.033 | 0.033 | 0.035 | 0.037 | 0.0696 | 0.105 | 0.034544 | 0.037 | 0.032842 | 0.0308 | 0.03 | 0.03 | 0.035 | 0.0369 | 0.0352 | 0.03 | 0.0354
| Mean |  | 0.031667 | 0.0385 | 0.05205 | 0.0342 | 0.040333 | 0.057283 | 0.034 | 0.0335 | 0.033333 | 0.035167 | 0.036833 | 0.06925 | 0.10983 | 0.035339 | 0.0375 | 0.033121 | 0.031417 | 0.03 | 0.03 | 0.034 | 0.040917 | 0.0351 | 0.031667 | 0.035683
| Median |  | 0.03 | 0.0385 | 0.05205 | 0.0343 | 0.0405 | 0.05755 | 0.034 | 0.0335 | 0.033 | 0.035 | 0.037 | 0.06945 | 0.107 | 0.035134 | 0.0375 | 0.033166 | 0.03155 | 0.03 | 0.03 | 0.034 | 0.041 | 0.03505 | 0.03 | 0.03575
| Std Dev. |  | 0.0040825 | 0.0015166 | 0.00018708 | 0.00032863 | 0.0008165 | 0.0010226 | 0 | 0.00054772 | 0.0010328 | 0.00040825 | 0.00040825 | 0.00057533 | 0.010759 | 0.00074448 | 0.0010488 | 0.00040072 | 0.00042622 | 0 | 0 | 0.00063246 | 0.0024547 | 0.00045607 | 0.0040825 | 0.00043551
| Rel.Std.Dev. |  | 12.89% | 3.939% | 0.3594% | 0.9609% | 2.024% | 1.785% | 0% | 1.635% | 3.098% | 1.161% | 1.108% | 0.8308% | 9.796% | 2.107% | 2.797% | 1.21% | 1.357% | 0% | 0% | 1.86% | 5.999% | 1.299% | 12.89% | 1.22%
| PDM3 |  | -8.417% | 11.35% | 50.53% | -1.09% | 16.65% | 65.67% | -1.669% | -3.115% | -3.597% | 1.705% | 6.525% | 100.3% | 217.6% | 2.205% | 8.454% | -4.211% | -9.14% | -13.24% | -13.24% | -1.669% | 18.33% | 1.512% | -8.417% | 3.2%
| Z-Score (Absolute) |  | 0.82 | 0.79 | 3.98 | 0.22 | 1.22 | 5.21 | 0.27 | 0.38 | 0.42 | 0.01 | 0.4 | 8.03 | 17.57 | 0.05 | 0.56 | 0.47 | 0.87 | 1.21 | 1.21 | 0.27 | 1.36 | 0.01 | 0.82 | 0.13
| Analytical results for Pb in OREAS 555b (Certified Value 13.6 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 39.83 | 11.82 | 14.7 | 14.5 | 8 | 14.4 | 13.8556 | 12.9 | 13.1 | 14.6 | 13.5 | 14.7 | 11.42 | 23 | 14.2987 | 15 | 14.7252 | 13 | 14.2 | 22 | 11.6 | 14.4 | 9 | 14.3 | 12.9 |  |  | 1
| 1 | 2 | 30.38 | 12.58 | 14.8 | 14.6 | 10 | 14.7 | 13.7996 | 13.8 | 12.8 | 14.9 | 13.1 | 14.2 | 10.99 | 43 | 14.0691 | 15 | 14.6973 | 12 | 13.6 | 20 | 12.6 | 14 | 9 | 13.5 | 13.2 |  |  | 11
| 1 | 3 | 33.29 | 11.3 | 14.3 | 14.3 | 12 | 14.9 | 14.2775 | 13.5 | 13.5 | 15.1 | 14 | 14.7 | 11.24 | 41 | 14.1069 | 15 | 14.8184 | 11 | 14 | 21 | 10.9 | 13.8 | 9 | 14.5 | 12.7 |  |  | 16
| 1 | 4 | 30.4 | 11.94 | 14.6 | 14.7 | 11 | 14.9 | 13.7332 | 13 | 12.5 | 14.2 | 13.5 | 13.9 | 11.36 | 35 | 14.111 | 15 | 14.7377 | 11 | 13.9 | 22 | 11.7 | 14 | 9 | 14 | 13.2 |  |  | 13.6067
| 1 | 5 | 32.52 | 11.71 | 15.8 | 14.5 | 12 | 14.9 | 14.0225 | 14 | 13.2 | 14.6 | 13 | 15.1 | 11.37 | 34 | 14.2631 | 15 | 14.6985 | 12 | 13.6 | 20 | 10.4 | 15.6 | 9 | 14.5 | 13 |  |  | 122
| 1 | 6 | 33.3 | 11.28 | 15.8 | 14.5 | 10 | 14.8 | 13.8674 | 12.6 | 13.1 | 14.6 | 13.6 | 14.7 | 11.28 | 32 | 14.2139 | 15 | 15.1607 | 13 | 13.8 | 22 | 12.5 | 16.1 | 9 | 14.1 | 13.1
| Mean |  | 33.2867 | 11.7717 | 15 | 14.5167 | 10.5 | 14.7667 | 13.926 | 13.3 | 13.0333 | 14.6667 | 13.45 | 14.55 | 11.2767 | 34.6667 | 14.1771 | 15 | 14.8063 | 12 | 13.85 | 21.1667 | 11.6167 | 14.65 | 9 | 14.15 | 13.0167
| Median |  | 32.905 | 11.765 | 14.75 | 14.5 | 10.5 | 14.85 | 13.8615 | 13.25 | 13.1 | 14.6 | 13.5 | 14.7 | 11.32 | 34.5 | 14.1624 | 15 | 14.7314 | 12 | 13.85 | 21.5 | 11.65 | 14.2 | 9 | 14.2 | 13.05
| Std Dev. |  | 3.4694 | 0.48002 | 0.64187 | 0.13292 | 1.5166 | 0.19664 | 0.1972 | 0.55136 | 0.34448 | 0.30768 | 0.36194 | 0.42778 | 0.15475 | 7.1181 | 0.094344 | 0 | 0.1792 | 0.89443 | 0.23452 | 0.98319 | 0.86583 | 0.96281 | 0 | 0.37815 | 0.19408
| Rel.Std.Dev. |  | 10.42% | 4.078% | 4.279% | 0.9156% | 14.44% | 1.332% | 1.416% | 4.146% | 2.643% | 2.098% | 2.691% | 2.94% | 1.372% | 20.53% | 0.6655% | 0% | 1.21% | 7.454% | 1.693% | 4.645% | 7.453% | 6.572% | 0% | 2.672% | 1.491%
| PDM3 |  | 144.6% | -13.49% | 10.24% | 6.688% | -22.83% | 8.525% | 2.346% | -2.254% | -4.214% | 7.79% | -1.152% | 6.933% | -17.12% | 154.8% | 4.192% | 10.24% | 8.816% | -11.81% | 1.788% | 55.56% | -14.63% | 7.668% | -33.86% | 3.993% | -4.336%
| Z-Score (Absolute) |  | 17.7 | 2.08 | 0.88 | 0.44 | NA | 0.67 | 0.1 | 0.68 | 0.92 | 0.58 | 0.54 | 0.47 | 2.54 | NA | 0.13 | NA | 0.71 | 1.87 | 0.17 | NA | 2.23 | 0.56 | NA | 0.1 | 0.94
Results from laboratories 5, 14, 17, 21 and 26 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 555b (Certified Value 16.9 ppm) | [28 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 1
2
| 1 | 1 | 14.1 | 17.61 | 18.5 | 17.7176 | 13 | 17.9277 | 16.9533 | 15.83 | 15.81 | 14.8 | 18.38 | [16 empty] | 1
| 1 | 2 | 12.3 | 18.11 | 18.4 | 17.6355 | 13 | 17.7977 | 17.057 | 15.99 | 15.26 | 17.1 | 18.35 | [16 empty] | 12
| 1 | 3 | 13.9 | 17.86 | 18.6 | 17.5933 | 13 | 18.007 | 17.0298 | 16.05 | 15.51 | 14.9 | 18.31 | [16 empty] | 16
| 1 | 4 | 14.4 | 18.01 | 17.3 | 17.6198 | 14 | 17.6336 | 18.1436 | 16.01 | 15.06 | 16.3 | 18.59 | [16 empty] | 16.9339
| 1 | 5 | 14.4 | 18.05 | 18.2 | 17.6193 | 14 | 17.5022 | 17.7649 | 15.85 | 15.42 | 18.1 | 18.55 | [16 empty] | 123
| 1 | 6 | 14.1 | 18.16 | 18.4 | 17.8114 | 14 | 17.4624 | 17.0586 | 15.63 | 16.31 | 16.7 | 18.46
| Mean |  | 13.8667 | 17.9667 | 18.2333 | 17.6662 | 13.5 | 17.7218 | 17.3345 | 15.8933 | 15.5617 | 16.3167 | 18.44
| Median |  | 14.1 | 18.03 | 18.4 | 17.6276 | 13.5 | 17.7157 | 17.0578 | 15.92 | 15.465 | 16.5 | 18.42
| Std Dev. |  | 0.79162 | 0.20265 | 0.4761 | 0.082849 | 0.54772 | 0.22488 | 0.49622 | 0.15667 | 0.44432 | 1.2844 | 0.11278
| Rel.Std.Dev. |  | 5.709% | 1.128% | 2.611% | 0.469% | 4.057% | 1.269% | 2.863% | 0.9858% | 2.855% | 7.872% | 0.6116%
| PDM3 |  | -18.11% | 6.099% | 7.674% | 4.324% | -20.28% | 4.653% | 2.366% | -6.145% | -8.103% | -3.645% | 8.894%
| Z-Score (Absolute) |  | 2.93 | 0.38 | 0.59 | 0.13 | NA | 0.18 | 0.13 | 1.3 | 1.56 | 0.95 | 0.76
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Pt in OREAS 555b (Indicative Value 10.5 ppb) | [28 empty] | Indicative
Pt | Round | Replicate | Lab | [26 empty] | 1
| No | No | 25 | [26 empty] | ppb
|  |  | 4A*OES/MS | [26 empty] | 1
1
| 1 | 1 | 17 | [26 empty] | 1
| 1 | 2 | 10 | [26 empty] | 22
| 1 | 3 | 8 | [26 empty] | 16
| 1 | 4 | 9 | [26 empty] | 10.5
| 1 | 5 | 10 | [26 empty] | 28
| 1 | 6 | 9
| Mean |  | 10.5
| Median |  | 9.5
| Std Dev. |  | 3.2711
| Rel.Std.Dev. |  | 31.15%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 555b (Certified Value 94 ppm) | [28 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 0
1
| 1 | 1 | 78 | 95.27 | 91.8 | 92.3 | 91.3645 | 90.6 | 86.8 | 95.4 | 98.1 | 103 | 99.2219 | 104 | 93.2092 | 94.5 | 92.3 | 97.3 | 84.9 | 83.9 | 98.56 | 94.8 | 97.6 | [6 empty] | 1
| 1 | 2 | 99.3 | 95.52 | 68.5 | 95.6 | 90.312 | 97.5 | 77.9 | 98.7 | 95.3 | 100.5 | 98.2991 | 105 | 93.0037 | 87.8 | 89.1 | 92.1 | 88.7 | 94.9 | 97.3 | 93.2 | 94.7 | [6 empty] | 13
| 1 | 3 | 55.5 | 97.37 | 83.8 | 94.6 | 88.8644 | 97.6 | 83.5 | 93.2 | 101 | 101 | 98.0236 | 106 | 91.8578 | 92.2 | 91.5 | 94.8 | 81.4 | 84 | 97.59 | 95.9 | 92.9 | [6 empty] | 16
| 1 | 4 | 79.1 | 98.15 | 99.1 | 97.4 | 92.8102 | 94.3 | 76.6 | 96.2 | 97.1 | 94.8 | 99.3867 | 106 | 95.1935 | 98 | 89.2 | 94.1 | 84.4 | 88.4 | 95.9 | 95.5 | 99.3 | [6 empty] | 94.4528
| 1 | 5 | 59.4 | 98.47 | 73.5 | 98.2 | 89.7122 | 96.4 | 86.3 | 90.1 | 92.6 | 100.5 | 100.0184 | 105 | 94.9896 | 93.2 | 91.6 | 92 | 83.7 | 96.9 | 98.44 | 97.2 | 92.7 | [6 empty] | 124
| 1 | 6 | 78.9 | 98.24 | 61.1 | 99.9 | 89.5179 | 94.6 | 78.7 | 88.2 | 97.9 | 99.9 | 98.0469 | 103 | 93.3823 | 92 | 90.8 | 97.2 | 88.1 | 91.1 | 96.28 | 94.9 | 102.4
| Mean |  | 75.0333 | 97.17 | 79.6333 | 96.3333 | 90.4302 | 95.1667 | 81.6333 | 93.6333 | 97 | 99.95 | 98.8328 | 104.8333 | 93.606 | 92.95 | 90.75 | 94.5833 | 85.2 | 89.8667 | 97.345 | 95.25 | 96.6
| Median |  | 78.45 | 97.76 | 78.65 | 96.5 | 90.0121 | 95.5 | 81.1 | 94.3 | 97.5 | 100.5 | 98.7605 | 105 | 93.2957 | 92.7 | 91.15 | 94.45 | 84.65 | 89.75 | 97.445 | 95.2 | 96.15
| Std Dev. |  | 15.8437 | 1.426 | 14.4872 | 2.7259 | 1.4388 | 2.6387 | 4.4684 | 3.9389 | 2.8383 | 2.7399 | 0.82719 | 1.169 | 1.2697 | 3.3453 | 1.3278 | 2.3387 | 2.7597 | 5.4496 | 1.0917 | 1.3278 | 3.8585
| Rel.Std.Dev. |  | 21.12% | 1.468% | 18.19% | 2.83% | 1.591% | 2.773% | 5.474% | 4.207% | 2.926% | 2.741% | 0.837% | 1.115% | 1.356% | 3.599% | 1.463% | 2.473% | 3.239% | 6.064% | 1.122% | 1.394% | 3.994%
| PDM3 |  | -20.56% | 2.877% | -15.69% | 1.991% | -4.259% | 0.7558% | -13.57% | -0.8676% | 2.697% | 5.82% | 4.637% | 10.99% | -0.8965% | -1.591% | -3.92% | 0.1382% | -9.796% | -4.855% | 3.062% | 0.844% | 2.273%
| Z-Score (Absolute) |  | 4.77 | 0.63 | 3.65 | 0.43 | 1.01 | 0.14 | 3.16 | 0.23 | 0.59 | 1.31 | 1.04 | 2.5 | 0.24 | 0.4 | 0.94 | 0 | 2.29 | 1.15 | 0.67 | 0.16 | 0.49
| Analytical results for Re in OREAS 555b (Certified Value 0.03 ppm) | [28 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 28 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | 0.029 | 0.025 | 0.03 | 0.033 | 0.031 | 0.022 | 0.022 | 0.032 | 0.033 | < 0.05 | < 0.05 | 0.025585 | 0.033 | 0.029 | 0.031 | 0.027 | 0.0289 | 0.034 | 0.02 | [8 empty] | 1
| 1 | 2 | 0.032 | 0.029 | 0.028 | 0.027 | 0.029 | 0.023 | 0.016 | 0.027 | 0.031 | < 0.05 | < 0.05 | 0.024732 | 0.031 | 0.032 | 0.032 | 0.026 | 0.0309 | 0.032 | 0.022 | [8 empty] | 14
| 1 | 3 | 0.026 | 0.029 | 0.028 | 0.032 | 0.03 | 0.022 | 0.025 | 0.033 | 0.034 | < 0.05 | < 0.05 | 0.024557 | 0.033 | 0.03 | 0.026 | 0.027 | 0.0277 | 0.033 | 0.022 | [8 empty] | 16
| 1 | 4 | 0.027 | 0.032 | 0.021 | 0.032 | 0.029 | 0.023 | 0.033 | 0.029 | 0.03 | < 0.05 | < 0.05 | 0.023241 | 0.033 | 0.035 | 0.029 | 0.026 | 0.0313 | 0.03 | 0.024 | [8 empty] | 0.029874
| 1 | 5 | 0.026 | 0.031 | 0.028 | 0.035 | 0.034 | 0.023 | 0.022 | 0.033 | 0.033 | < 0.05 | < 0.05 | 0.023769 | 0.034 | 0.03 | 0.029 | 0.024 | 0.0331 | 0.033 | 0.022 | [8 empty] | 125
| 1 | 6 | 0.026 | 0.035 | 0.029 | 0.036 | 0.033 | 0.021 | 0.024 | 0.035 | 0.029 | < 0.05 | < 0.05 | 0.024045 | 0.031 | 0.032 | 0.028 | 0.027 | 0.0322 | 0.029 | 0.023
| Mean |  | 0.027667 | 0.030167 | 0.027333 | 0.0325 | 0.031 | 0.022333 | 0.023667 | 0.0315 | 0.031667 |  |  | 0.024322 | 0.0325 | 0.031333 | 0.029167 | 0.026167 | 0.030683 | 0.031833 | 0.022167
| Median |  | 0.0265 | 0.03 | 0.028 | 0.0325 | 0.0305 | 0.0225 | 0.023 | 0.0325 | 0.032 |  |  | 0.024301 | 0.033 | 0.031 | 0.029 | 0.0265 | 0.0311 | 0.0325 | 0.022
| Std Dev. |  | 0.0024221 | 0.0033714 | 0.0032042 | 0.0031464 | 0.0020976 | 0.0008165 | 0.0055377 | 0.0029496 | 0.0019664 |  |  | 0.00082105 | 0.0012247 | 0.0021602 | 0.002137 | 0.001169 | 0.0020322 | 0.0019408 | 0.0013292
| Rel.Std.Dev. |  | 8.755% | 11.18% | 11.72% | 9.681% | 6.767% | 3.656% | 23.4% | 9.364% | 6.21% |  |  | 3.376% | 3.768% | 6.894% | 7.327% | 4.468% | 6.623% | 6.097% | 5.996%
| PDM3 |  | -7.389% | 0.9791% | -8.505% | 8.79% | 3.769% | -25.24% | -20.78% | 5.442% | 6% |  |  | -18.59% | 8.79% | 4.884% | -2.368% | -12.41% | 2.709% | 6.558% | -25.8%
| Z-Score (Absolute) |  | 0.38 | 0.25 | 0.46 | 0.84 | 0.46 | 1.73 | 1.39 | 0.59 | 0.63 | 1.05 | 1.05 | 1.23 | 0.84 | 0.55 | 0 | 0.76 | 0.38 | 0.67 | 1.77
| Analytical results for S in OREAS 555b (Certified Value 4.39 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.76 | 4.68 | 4.56 | 4.441 | 4.7 | 3.37 | 4.28 | 4.21 | 4.35 | 4.33 | 4.52 | 4.55 | 4.3044 | 4.2447 | 4.35 | 4.6878 | 3.8449 | 4.38 | 4.41 | 4.28 | 4.167 | 3.86 | 4.4454 | 4.23 | 4.45 |  |  | 1
| 1 | 2 | 3.56 | 4.26 | 4.56 | 4.3959 | 4.66 | 3.41 | 4.25 | 4.22 | 4.54 | 4.38 | 4.48 | 4.52 | 4.3932 | 4.2145 | 4.39 | 4.7087 | 3.8743 | 4.34 | 4.49 | 4.35 | 4.288 | 3.91 | 4.486 | 4.23 | 4.46 |  |  | 28
| 1 | 3 | 3.5 | 4.58 | 4.59 | 4.3906 | 4.52 | 3.45 | 4.35 | 4.24 | 4.55 | 4.5 | 4.52 | 4.64 | 4.4207 | 4.2596 | 4.41 | 4.6587 | 3.8943 | 4.33 | 4.41 | 4.31 | 4.211 | 4.07 | 4.4338 | 4.2 | 4.35 |  |  | 16
| 1 | 4 | 3.45 | 4.78 | 4.62 | 4.3275 | 4.54 | 3.48 | 4.28 | 4.27 | 4.52 | 4.35 | 4.52 | 4.59 | 4.3222 | 4.2261 | 4.5 | 4.6771 | 3.8076 | 4.35 | 4.44 | 4.25 | 4.027 | 4.19 | 4.3956 | 4.15 | 4.51 |  |  | 4.3937
| 1 | 5 | 3.49 | 4.8 | 4.61 | 4.4027 | 4.48 | 3.47 | 4.27 | 4.25 | 4.48 | 4.32 | 4.54 | 4.52 | 4.4497 | 4.2894 | 4.32 | 4.647 | 3.8468 | 4.25 | 4.42 | 4.24 | 4.149 | 4.07 | 4.527 | 4.25 | 4.47 |  |  | 126
| 1 | 6 | 3.24 | 4.63 | 4.6 | 4.3703 | 4.5 | 3.47 | 4.24 | 4.21 | 4.48 | 4.38 | 4.53 | 4.48 | 4.2847 | 4.2169 | 4.48 | 4.6534 | 3.7506 | 4.24 | 4.43 | 4.32 | 3.992 | 4.29 | 4.4448 | 4.21 | 4.4
| Mean |  | 3.5 | 4.6217 | 4.59 | 4.388 | 4.5667 | 3.4417 | 4.2783 | 4.2333 | 4.4867 | 4.3767 | 4.5183 | 4.55 | 4.3625 | 4.2419 | 4.4083 | 4.6721 | 3.8364 | 4.315 | 4.4333 | 4.2917 | 4.139 | 4.065 | 4.4554 | 4.2117 | 4.44
| Median |  | 3.495 | 4.655 | 4.595 | 4.3933 | 4.53 | 3.46 | 4.275 | 4.23 | 4.5 | 4.365 | 4.52 | 4.535 | 4.3577 | 4.2354 | 4.4 | 4.6679 | 3.8458 | 4.335 | 4.425 | 4.295 | 4.158 | 4.07 | 4.4451 | 4.22 | 4.455
| Std Dev. |  | 0.16817 | 0.19641 | 0.025298 | 0.037597 | 0.090921 | 0.043089 | 0.038687 | 0.024221 | 0.073121 | 0.06532 | 0.020412 | 0.057271 | 0.067785 | 0.028992 | 0.070828 | 0.023542 | 0.051304 | 0.056833 | 0.030111 | 0.042622 | 0.11173 | 0.1627 | 0.045427 | 0.034881 | 0.056569
| Rel.Std.Dev. |  | 4.805% | 4.25% | 0.5512% | 0.8568% | 1.991% | 1.252% | 0.9042% | 0.5722% | 1.63% | 1.492% | 0.4518% | 1.259% | 1.554% | 0.6835% | 1.607% | 0.5039% | 1.337% | 1.317% | 0.6792% | 0.9931% | 2.7% | 4.002% | 1.02% | 0.8282% | 1.274%
| PDM3 |  | -20.34% | 5.188% | 4.467% | -0.1303% | 3.936% | -21.67% | -2.626% | -3.65% | 2.115% | -0.3882% | 2.836% | 3.557% | -0.7109% | -3.456% | 0.3325% | 6.336% | -12.68% | -1.792% | 0.9015% | -2.323% | -5.797% | -7.482% | 1.405% | -4.144% | 1.053%
| Z-Score (Absolute) |  | 4.17 | 1.17 | 1.02 | 0.05 | 0.9 | 4.45 | 0.47 | 0.68 | 0.52 | 0 | 0.67 | 0.83 | 0.07 | 0.64 | 0.15 | 1.41 | 2.57 | 0.29 | 0.27 | 0.4 | 1.13 | 1.48 | 0.37 | 0.79 | 0.3
| Analytical results for Sb in OREAS 555b (Certified Value 3.72 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 5 | 3.96 | < 5 | 3.7 | 4.4 | 3.72 | 3.5924 | 3.35 | 3.47 | 3.38 | 3.65 | 3.9 | < 5 | 4.0642 | 3.5 | 3.7755 | NA | 3.54 | 7 | 4.3 | 3.53 | 3.4 | 4.03 | 2.9 | 3.7 |  |  | 1
| 1 | 2 | < 5 | 4.05 | < 5 | 3.65 | 4.3 | 3.76 | 4.1268 | 3.44 | 3.39 | 3.82 | 3.59 | 4.02 | < 5 | 3.7596 | 3.5 | 3.7332 | NA | 3.56 | 6 | 4.1 | 3.33 | 3.3 | 3.86 | 2.8 | 3.8 |  |  | 15
| 1 | 3 | < 5 | 3.9 | < 5 | 3.86 | 4.2 | 3.86 | 4.0087 | 3.43 | 3.42 | 3.49 | 3.69 | 3.98 | < 5 | 3.8108 | 3.4 | 3.7072 | NA | 3.83 | 8 | 3.6 | 3.45 | 3.3 | 3.99 | 2.9 | 3.8 |  |  | 16
| 1 | 4 | < 5 | 3.95 | < 5 | 3.66 | 3.2 | 3.86 | 3.5984 | 3.5 | 3.38 | 3.76 | 4.01 | 3.55 | < 5 | 3.9492 | 3.6 | 3.7732 | NA | 3.55 | 6 | 3.8 | 3.25 | 3.3 | 3.9 | 2.9 | 4.2 |  |  | 3.7211
| 1 | 5 | < 5 | 4.02 | < 5 | 3.85 | 4 | 3.76 | 3.5329 | 3.49 | 3.56 | 3.7 | 3.47 | 3.87 | < 5 | 3.9788 | 3.3 | 3.7888 | NA | 3.79 | 7 | 4 | 3.41 | 3.7 | 4.02 | 2.9 | 3.8 |  |  | 127
| 1 | 6 | < 5 | 4.11 | < 5 | 3.89 | 4 | 3.73 | 3.5209 | 3.51 | 3.46 | 3.61 | 3.63 | 3.68 | < 5 | 4.0431 | 3.5 | 3.6364 | NA | 3.55 | 6 | 4.1 | 3.52 | 3.5 | 3.88 | 2.9 | 3.8
| Mean |  |  | 3.9983 |  | 3.7683 | 4.0167 | 3.7817 | 3.73 | 3.4533 | 3.4467 | 3.6267 | 3.6733 | 3.8333 |  | 3.9343 | 3.4667 | 3.7357 |  | 3.6367 | 6.6667 | 3.9833 | 3.415 | 3.4167 | 3.9467 | 2.8833 | 3.85
| Median |  |  | 3.99 |  | 3.775 | 4.1 | 3.76 | 3.5954 | 3.465 | 3.44 | 3.655 | 3.64 | 3.885 |  | 3.964 | 3.5 | 3.7532 |  | 3.555 | 6.5 | 4.05 | 3.43 | 3.35 | 3.945 | 2.9 | 3.8
| Std Dev. |  |  | 0.076267 |  | 0.1098 | 0.43089 | 0.062743 | 0.26604 | 0.060222 | 0.066232 | 0.16753 | 0.18129 | 0.18217 |  | 0.12384 | 0.10328 | 0.057416 |  | 0.13501 | 0.8165 | 0.24833 | 0.1095 | 0.16021 | 0.075277 | 0.040825 | 0.17607
| Rel.Std.Dev. |  |  | 1.907% |  | 2.914% | 10.73% | 1.659% | 7.132% | 1.744% | 1.922% | 4.619% | 4.935% | 4.752% |  | 3.148% | 2.979% | 1.537% |  | 3.712% | 12.25% | 6.234% | 3.206% | 4.689% | 1.907% | 1.416% | 4.573%
| PDM3 |  |  | 7.451% |  | 1.27% | 7.943% | 1.628% | 0.2399% | -7.196% | -7.375% | -2.537% | -1.283% | 3.016% |  | 5.729% | -6.837% | 0.3934% |  | -2.269% | 79.16% | 7.048% | -8.226% | -8.181% | 6.062% | -22.51% | 3.464%
| Z-Score (Absolute) |  | 3.49 | 0.97 | 3.49 | 0.28 | 1.02 | 0.32 | 0.17 | 0.65 | 0.67 | 0.14 | 0 | 0.48 | 3.49 | 0.78 | 0.61 | 0.19 | NA | 0.11 | NA | 0.92 | 0.77 | 0.76 | 0.81 | 2.35 | 0.53
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 555b (Certified Value 12.6 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 12.9 | 12.4 | 12.7 | 13 | 12.8 | 9.1759 | 12 | 11.8 | 12.3 | 13.2 | 13 | 12 | 14.4989 | 11.8135 | 12 | 12.6252 | 11.98 | 12.4 | 11.6 | 14.3 | 10 | 12 | 13.7 | 12 | 11.2 |  |  | 1
| 1 | 2 | 12.5 | 12.7 | 12.8 | 13 | 13.4 | 9.2724 | 12.8 | 11.5 | 12.8 | 12.8 | 12.7 | 11 | 14.5252 | 11.4833 | 12 | 12.8635 | 11.97 | 12 | 12.1 | 14.6 | 10 | 12 | 12.6 | 12 | 11.4 |  |  | 1
| 1 | 3 | 12.9 | 12.1 | 12.8 | 12 | 13.2 | 9.4119 | 12.6 | 12.5 | 12.2 | 12.9 | 12.8 | 12 | 14.3387 | 11.376 | 12 | 12.7459 | 12.22 | 12.4 | 11.5 | 13.9 | 9.5 | 12 | 12.7 | 12 | 11.1 |  |  | 16
| 1 | 4 | 13.3 | 12.9 | 13.1 | 12 | 13.4 | 8.8628 | 12.5 | 11.7 | 12.8 | 13 | 12.1 | 12 | 14.3679 | 11.5823 | 12 | 13.1343 | 11.88 | 12.3 | 12 | 14.8 | 9.7 | 13 | 13 | 12 | 11.6 |  |  | 12.5862
| 1 | 5 | 13.5 | 12.4 | 13.1 | 12 | 13.9 | 9.0592 | 12.8 | 12.2 | 12.4 | 12 | 13 | 12 | 14.2745 | 11.2947 | 12 | 13.1407 | 12.02 | 12.4 | 11.6 | 14.2 | 9.7 | 12 | 13.1 | 12 | 11.2 |  |  | 128
| 1 | 6 | 13 | 12.7 | 13 | 12 | 14.1 | 8.8036 | 12.4 | 12 | 12 | 12.7 | 12.8 | 11 | 14.2959 | 11.5426 | 12 | 13.0526 | 11.87 | 12.6 | 12.1 | 15.5 | 10.1 | 14 | 12.8 | 12 | 11.5
| Mean |  | 13.0167 | 12.5333 | 12.9167 | 12.3333 | 13.4667 | 9.0976 | 12.5167 | 11.95 | 12.4167 | 12.7667 | 12.7333 | 11.6667 | 14.3835 | 11.5154 | 12 | 12.9271 | 11.99 | 12.35 | 11.8167 | 14.55 | 9.8333 | 12.5 | 12.9833 | 12 | 11.3333
| Median |  | 12.95 | 12.55 | 12.9 | 12 | 13.4 | 9.1175 | 12.55 | 11.9 | 12.35 | 12.85 | 12.8 | 12 | 14.3533 | 11.513 | 12 | 12.9581 | 11.975 | 12.4 | 11.8 | 14.45 | 9.85 | 12 | 12.9 | 12 | 11.3
| Std Dev. |  | 0.34881 | 0.28752 | 0.17224 | 0.5164 | 0.47188 | 0.23599 | 0.29944 | 0.36194 | 0.32506 | 0.41312 | 0.33267 | 0.5164 | 0.10508 | 0.18071 | 0 | 0.21557 | 0.12712 | 0.19748 | 0.27869 | 0.56125 | 0.23381 | 0.83666 | 0.39707 | 0 | 0.19664
| Rel.Std.Dev. |  | 2.68% | 2.294% | 1.333% | 4.187% | 3.504% | 2.594% | 2.392% | 3.029% | 2.618% | 3.236% | 2.613% | 4.426% | 0.7306% | 1.569% | 0% | 1.668% | 1.06% | 1.599% | 2.358% | 3.857% | 2.378% | 6.693% | 3.058% | 0% | 1.735%
| PDM3 |  | 3.42% | -0.4203% | 2.625% | -2.009% | 6.995% | -27.72% | -0.5527% | -5.055% | -1.347% | 1.434% | 1.169% | -7.306% | 14.28% | -8.508% | -4.658% | 2.708% | -4.737% | -1.877% | -6.114% | 15.6% | -21.87% | -0.6851% | 3.155% | -4.658% | -9.955%
| Z-Score (Absolute) |  | 0.65 | 0.01 | 0.51 | NA | 1.24 | 4.5 | 0.01 | 0.76 | 0.14 | 0.32 | 0.27 | NA | 2.44 | 1.33 | NA | 0.53 | 0.7 | 0.23 | 0.93 | 2.66 | 3.54 | NA | 0.6 | NA | 1.57
Results from laboratories 5, 14, 17, 25 and 27 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 555b (Certified Value 6.97 ppm) | [28 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 25 | 26 | 27 | 28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | 6.3 | 8.27 | 6.5 | 6.9 | 6.2 | 8 | 7 | 8 | 6 | 9 | < 10 | 6.6677 | 7 | 7.2423 | NA | 5.8 | 5.53 | 6.7 | 8 | 7 | [7 empty] | 1
| 1 | 2 | 5.8 | 8.72 | 8 | 6.5 | 6 | 8 | 7 | 8 | 7 | 9 | < 10 | 8.2072 | 7 | 7.1624 | NA | 5.7 | 6.13 | 6.6 | 8 | 7 | [7 empty] | 19
| 1 | 3 | 6.6 | 8.39 | 7 | 6.8 | 6.1 | 8 | 6 | 7 | 7 | 8 | < 10 | 5.7371 | 7 | 7.1473 | NA | 5.6 | 5.24 | 6.5 | 6 | 7 | [7 empty] | 16
| 1 | 4 | 6.1 | 8.61 | 6.5 | 4.7 | 6.4 | 7 | 6 | 8 | 7 | 7 | < 10 | 7.0513 | 7 | 7.2697 | NA | 5.8 | 6.09 | 6.5 | 7 | 8 | [7 empty] | 6.9708
| 1 | 5 | 6.4 | 8.68 | 8.2 | 6.4 | 6 | 8 | 7 | 9 | 7 | 8 | < 10 | 6.3198 | 7 | 7.7291 | NA | 5.8 | 5.84 | 6.3 | 7 | 7 | [7 empty] | 129
| 1 | 6 | 6.2 | 8.41 | 7.5 | 6.4 | 6.1 | 8 | 6 | 7 | 7 | 7 | < 10 | 7.9508 | 7 | 7.2203 | NA | 5.9 | 5.15 | 6.9 | 8 | 7
| Mean |  | 6.2333 | 8.5133 | 7.2833 | 6.2833 | 6.1333 | 7.8333 | 6.5 | 7.8333 | 6.8333 | 8 |  | 6.989 | 7 | 7.2952 |  | 5.7667 | 5.6633 | 6.5833 | 7.3333 | 7.1667
| Median |  | 6.25 | 8.51 | 7.25 | 6.45 | 6.1 | 8 | 6.5 | 8 | 7 | 8 |  | 6.8595 | 7 | 7.2313 |  | 5.8 | 5.685 | 6.55 | 7.5 | 7
| Std Dev. |  | 0.27325 | 0.18162 | 0.73598 | 0.80353 | 0.15055 | 0.40825 | 0.54772 | 0.75277 | 0.40825 | 0.89443 |  | 0.95198 | 0 | 0.21764 |  | 0.10328 | 0.42226 | 0.20412 | 0.8165 | 0.40825
| Rel.Std.Dev. |  | 4.384% | 2.133% | 10.1% | 12.79% | 2.455% | 5.212% | 8.427% | 9.61% | 5.974% | 11.18% |  | 13.62% | 0% | 2.983% |  | 1.791% | 7.456% | 3.101% | 11.13% | 5.696%
| PDM3 |  | -10.58% | 22.13% | 4.484% | -9.862% | -12.01% | 12.37% | -6.754% | 12.37% | -1.972% | 14.76% |  | 0.261% | 0.4192% | 4.654% |  | -17.27% | -18.76% | -5.558% | 5.201% | 2.81%
| Z-Score (Absolute) |  | 0.72 | 1.46 | 0.28 | 0.67 | 0.82 | 0.81 | 0.47 | 0.81 | 0.15 | 0.97 | 1.9 | 0 | 0.01 | 0.29 | NA | 1.17 | 1.27 | 0.39 | 0.33 | 0.17
| Analytical results for Sm in OREAS 555b (Certified Value 8.87 ppm) | [28 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 23 | 25 | 26 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [16 empty] | 2
3
| 1 | 1 | 6.1 | 8.91 | 10 | 8.8581 | 7 | 9.5411 | 8.7108 | 8.4 | 8.25 | 7.88 | 9.35 | [16 empty] | 1
| 1 | 2 | 6.4 | 8.88 | 9.8 | 9.0803 | 7 | 9.2539 | 8.5715 | 7.84 | 7.99 | 9.11 | 9.2 | [16 empty] | 3
| 1 | 3 | 5.8 | 8.94 | 10.3 | 8.9203 | 7 | 9.1403 | 8.5954 | 8.15 | 8.04 | 7.85 | 9.3 | [16 empty] | 16
| 1 | 4 | 6.2 | 9.05 | 7.6 | 8.5562 | 7 | 9.1789 | 8.9262 | 8.74 | 7.78 | 8.26 | 9.39 | [16 empty] | 8.8742
| 1 | 5 | 5.9 | 8.89 | 9.5 | 8.8232 | 7 | 9.2281 | 8.8886 | 8.23 | 7.92 | 9.44 | 9.24 | [16 empty] | 130
| 1 | 6 | 5.8 | 9.17 | 10 | 8.8247 | 7 | 9.0115 | 8.9295 | 8.09 | 8.33 | 8.61 | 9.42
| Mean |  | 6.0333 | 8.9733 | 9.5333 | 8.8438 | 7 | 9.2256 | 8.7703 | 8.2417 | 8.0517 | 8.525 | 9.3167
| Median |  | 6 | 8.925 | 9.9 | 8.8414 | 7 | 9.2035 | 8.7997 | 8.19 | 8.015 | 8.435 | 9.325
| Std Dev. |  | 0.24221 | 0.11431 | 0.98319 | 0.1706 | 0 | 0.17643 | 0.16569 | 0.30538 | 0.20586 | 0.65246 | 0.085946
| Rel.Std.Dev. |  | 4.015% | 1.274% | 10.31% | 1.929% | 0% | 1.912% | 1.889% | 3.705% | 2.557% | 7.654% | 0.9225%
| PDM3 |  | -32.01% | 1.117% | 7.427% | -0.3431% | -21.12% | 3.96% | -1.171% | -7.128% | -9.269% | -3.935% | 4.986%
| Z-Score (Absolute) |  | 4.03 | 0.24 | 1.06 | 0.05 | NA | 0.61 | 0.05 | 0.82 | 1.1 | 0.41 | 0.74
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 555b (Certified Value 2.99 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | < 5 | 3.4 | < 10 | 3.1 | 3 | 3.6 | 3.0598 | 2.5 | 2.7 | 2.9 | 3.1 | 3.3 | < 10 | < 1 | 3.0925 | 3 | 3.0258 | 2.87 | 3.2 | < 10 | 3.7 | 2.7 | 2.6 | 3.1 | 2.7 | 2.8 |  | 1
| 1 | 2 | < 5 | 3.6 | < 10 | 2.8 | 3 | 3.6 | 3.0593 | 2.6 | 2.7 | 3 | 3 | 3.1 | < 10 | < 1 | 3.1274 | 3.1 | 3.0238 | 2.57 | 3 | < 10 | 3.4 | 2.5 | 3 | 3.2 | 2.6 | 2.9 |  | 20
| 1 | 3 | < 5 | 3.3 | < 10 | 2.9 | 3 | 3.7 | 2.9672 | 2.7 | 2.7 | 2.9 | 3.1 | 3.2 | < 10 | < 1 | 3.2963 | 3.1 | 2.9071 | 1.97 | 3.2 | < 10 | 3.3 | 2.7 | 2.6 | 3 | 2.8 | 2.9 |  | 16
| 1 | 4 | < 5 | 3.3 | < 10 | 3.2 | 2 | 3.6 | 3.0446 | 2.6 | 2.5 | 3 | 3 | 3 | < 10 | < 1 | 3.136 | 3 | 3.0199 | 2.47 | 3.1 | < 10 | 3.4 | 2.5 | 2.8 | 3.2 | 2.7 | 3.1 |  | 2.9872
| 1 | 5 | < 5 | 3.3 | < 10 | 3.1 | 3 | 3.7 | 3.0479 | 2.7 | 2.8 | 3 | 2.9 | 3.1 | < 10 | < 1 | 3.2981 | 3 | 3.0692 | 2.37 | 3.2 | < 10 | 3.3 | 2.5 | 3 | 3.1 | 2.7 | 2.9 |  | 131
| 1 | 6 | < 5 | 3.4 | < 10 | 3.1 | 3 | 3.8 | 3.009 | 2.6 | 2.6 | 3 | 2.9 | 3.1 | < 10 | < 1 | 3.1505 | 3.1 | 3.0318 | 1.98 | 3.2 | < 10 | 3.8 | 2.7 | 2.8 | 3 | 2.7 | 2.8
| Mean |  |  | 3.3833 |  | 3.0333 | 2.8333 | 3.6667 | 3.0313 | 2.6167 | 2.6667 | 2.9667 | 3 | 3.1333 |  |  | 3.1835 | 3.05 | 3.013 | 2.3717 | 3.15 |  | 3.4833 | 2.6 | 2.8 | 3.1 | 2.7 | 2.9
| Median |  |  | 3.35 |  | 3.1 | 3 | 3.65 | 3.0463 | 2.6 | 2.7 | 3 | 3 | 3.1 |  |  | 3.1432 | 3.05 | 3.0248 | 2.42 | 3.2 |  | 3.4 | 2.6 | 2.8 | 3.1 | 2.7 | 2.9
| Std Dev. |  |  | 0.1169 |  | 0.15055 | 0.40825 | 0.08165 | 0.036489 | 0.075277 | 0.10328 | 0.05164 | 0.089443 | 0.10328 |  |  | 0.090125 | 0.054772 | 0.054863 | 0.34988 | 0.083666 |  | 0.2137 | 0.10954 | 0.17889 | 0.089443 | 0.063246 | 0.10954
| Rel.Std.Dev. |  |  | 3.455% |  | 4.963% | 14.41% | 2.227% | 1.204% | 2.877% | 3.873% | 1.741% | 2.981% | 3.296% |  |  | 2.831% | 1.796% | 1.821% | 14.75% | 2.656% |  | 6.135% | 4.213% | 6.389% | 2.885% | 2.342% | 3.777%
| PDM3 |  |  | 13.26% |  | 1.546% | -5.15% | 22.75% | 1.478% | -12.4% | -10.73% | -0.6861% | 0.4298% | 4.893% |  |  | 6.571% | 2.104% | 0.8634% | -20.6% | 5.451% |  | 16.61% | -12.96% | -6.265% | 3.777% | -9.613% | -2.918%
| Z-Score (Absolute) |  | 1.08 | 0.72 | 4.01 | 0 | NA | 1.29 | 0 | 0.84 | 0.74 | 0.13 | 0.06 | 0.21 | 4.01 | 5.15 | 0.31 | 0.04 | 0.04 | 1.34 | 0.24 | 4.01 | 0.92 | 0.88 | 0.47 | 0.14 | 0.67 | 0.27
| Analytical results for Sr in OREAS 555b (Certified Value 70 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 69.18 | 75 | 68.9 | 72.7 | 73.8 | 64.7 | 65.9928 | 69.7 | 67.6 | 73 | 71.8 | 71.3 | 59.61 | 65 | 73.6302 | < 100 | 71.5 | 68.4467 | 75 | 68.4 | 65 | 73 | 60 | 54 | 72.64 | 68 | 68.6 | 1
| 1 | 2 | 67.23 | 69 | 68.6 | 72.72 | 71.3 | 66.8 | 67.0022 | 72.8 | 66.9 | 75.4 | 70.4 | 69.4 | 58.41 | 65 | 72.9211 | < 100 | 72.5 | 69.4558 | 76 | 64.9 | 66.4 | 74 | 61.3 | 51.6 | 71.31 | 66 | 69.3 | 32
| 1 | 3 | 70.34 | 73 | 68.1 | 73.3 | 71.4 | 66.4 | 66.6046 | 72.1 | 70.2 | 73.3 | 72.9 | 71.9 | 59.52 | 65 | 73.3635 | < 100 | 73.5 | 68.2683 | 74 | 66.9 | 65.1 | 74 | 61.8 | 57.8 | 72.25 | 68 | 70.2 | 16
| 1 | 4 | 68.53 | 77 | 68.4 | 73.83 | 71 | 68.8 | 67.7068 | 72.5 | 66 | 73.8 | 70.9 | 66.8 | 59.27 | 65 | 72.955 | < 100 | 73.5 | 70.1969 | 77 | 66.9 | 67 | 74 | 59.3 | 61.2 | 71.99 | 67 | 73.3 | 69.8607
| 1 | 5 | 65.23 | 76 | 68.8 | 75.19 | 71.6 | 69.7 | 67.3996 | 72.9 | 70.1 | 73.2 | 67.6 | 70.1 | 59.32 | 64 | 73.5863 | < 100 | 74.3 | 69.9623 | 76 | 67.3 | 65.4 | 76 | 63 | 57.8 | 72.67 | 68 | 69.7 | 132
| 1 | 6 | 67.23 | 73 | 68.1 | 75.2 | 70.9 | 69.5 | 67.192 | 71 | 68.2 | 72.3 | 70.4 | 69.1 | 59.46 | 65 | 72.6224 | < 100 | 72 | 69.1825 | 74 | 66.9 | 67.2 | 74 | 61.5 | 61.7 | 71.35 | 68 | 69.8
| Mean |  | 67.9567 | 73.8333 | 68.4833 | 73.8233 | 71.6667 | 67.65 | 66.983 | 71.8333 | 68.1667 | 73.5 | 70.6667 | 69.7667 | 59.265 | 64.8333 | 73.1798 |  | 72.8833 | 69.2521 | 75.3333 | 66.8833 | 66.0167 | 74.1667 | 61.15 | 57.35 | 72.035 | 67.5 | 70.15
| Median |  | 67.88 | 74 | 68.5 | 73.565 | 71.35 | 67.8 | 67.0971 | 72.3 | 67.9 | 73.25 | 70.65 | 69.75 | 59.39 | 65 | 73.1593 |  | 73 | 69.3192 | 75.5 | 66.9 | 65.9 | 74 | 61.4 | 57.8 | 72.12 | 68 | 69.75
| Std Dev. |  | 1.7886 | 2.8577 | 0.34303 | 1.1416 | 1.0764 | 1.9967 | 0.61076 | 1.2517 | 1.7025 | 1.0507 | 1.7818 | 1.8107 | 0.4373 | 0.40825 | 0.40741 |  | 1.0591 | 0.7823 | 1.2111 | 1.1321 | 0.97656 | 0.98319 | 1.3217 | 3.9607 | 0.60205 | 0.83666 | 1.6355
| Rel.Std.Dev. |  | 2.632% | 3.871% | 0.5009% | 1.546% | 1.502% | 2.952% | 0.9118% | 1.742% | 2.498% | 1.43% | 2.521% | 2.595% | 0.7379% | 0.6297% | 0.5567% |  | 1.453% | 1.13% | 1.608% | 1.693% | 1.479% | 1.326% | 2.161% | 6.906% | 0.8358% | 1.239% | 2.331%
| PDM3 |  | -2.725% | 5.687% | -1.972% | 5.672% | 2.585% | -3.164% | -4.119% | 2.824% | -2.425% | 5.209% | 1.154% | -0.1346% | -15.17% | -7.196% | 4.751% |  | 4.327% | -0.8711% | 7.834% | -4.262% | -5.502% | 6.164% | -12.47% | -17.91% | 3.112% | -3.379% | 0.4141%
| Z-Score (Absolute) |  | 0.31 | 1.11 | 0.19 | 1.11 | 0.59 | 0.39 | 0.55 | 0.63 | 0.26 | 1.03 | 0.34 | 0.12 | 2.42 | 1.07 | 0.95 | 4.66 | 0.88 | 0 | 1.47 | 0.57 | 0.78 | 1.19 | 1.96 | 2.88 | 0.67 | 0.42 | 0.22
| Analytical results for Ta in OREAS 555b (Certified Value 1.19 ppm) | [28 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.6 | < 5 | 1.03 | 1.9 | 0.73 | 1.8331 | 1.2 | 1.2 | 1.37 | 1.24 | 1.49 | < 1 | 1.3999 | 0.41 | 1.1099 | NA | 1.5 | 1.2 | 0.79 | 1.05 | 1.15 | 1 | 1.2 |  |  |  |  | 1
| 1 | 2 | 1.6 | < 5 | 1.09 | 1.9 | 0.71 | 1.8322 | 1.24 | 1.17 | 1.32 | 1.18 | 1.46 | < 1 | 1.3659 | 0.45 | 1.0639 | NA | 1.5 | 1.1 | 0.82 | 1.09 | 1.18 | 0.9 | 1.3 |  |  |  |  | 5
| 1 | 3 | 1.5 | < 5 | 1.06 | 1.9 | 0.75 | 1.8429 | 1.28 | 1.24 | 1.38 | 1.19 | 1.51 | < 1 | 1.4184 | 0.55 | 1.0315 | NA | 1.6 | 1.2 | 0.79 | 0.96 | 1.21 | 1 | 1.18 |  |  |  |  | 16
| 1 | 4 | 1.5 | < 5 | 0.93 | 1.4 | 0.68 | 1.8304 | 1.21 | 1.19 | 1.35 | 1.19 | 1.41 | < 1 | 1.4128 | 0.53 | 1.0477 | NA | 1.5 | 1.1 | 0.77 | 1.15 | 1.21 | 1 | 1.25 |  |  |  |  | 1.192
| 1 | 5 | 1.6 | < 5 | 1.08 | 1.8 | 0.65 | 1.8188 | 1.28 | 1.21 | 1.42 | 1.08 | 1.49 | < 1 | 1.3935 | 0.5 | 1.0717 | NA | 1.6 | 1.2 | 0.82 | 1.23 | 1.23 | 1 | 1.22 |  |  |  |  | 133
| 1 | 6 | 1.6 | < 5 | 0.94 | 1.8 | 0.77 | 1.842 | 1.28 | 1.15 | 1.4 | 1.22 | 1.44 | < 1 | 1.3662 | 0.5 | 1.072 | NA | 1.6 | 1.1 | 0.87 | 1.14 | 1.18 | 1 | 1.21
| Mean |  | 1.5667 |  | 1.0217 | 1.7833 | 0.715 | 1.8332 | 1.2483 | 1.1933 | 1.3733 | 1.1833 | 1.4667 |  | 1.3928 | 0.49 | 1.0661 |  | 1.55 | 1.15 | 0.81 | 1.1033 | 1.1933 | 0.98333 | 1.2267
| Median |  | 1.6 |  | 1.045 | 1.85 | 0.72 | 1.8326 | 1.26 | 1.195 | 1.375 | 1.19 | 1.475 |  | 1.3967 | 0.5 | 1.0678 |  | 1.55 | 1.15 | 0.805 | 1.115 | 1.195 | 1 | 1.215
| Std Dev. |  | 0.05164 |  | 0.070261 | 0.19408 | 0.044609 | 0.0088109 | 0.037103 | 0.031411 | 0.03559 | 0.055377 | 0.037238 |  | 0.02253 | 0.051769 | 0.026581 |  | 0.054772 | 0.054772 | 0.035214 | 0.09288 | 0.028752 | 0.040825 | 0.04274
| Rel.Std.Dev. |  | 3.296% |  | 6.877% | 10.88% | 6.239% | 0.4806% | 2.972% | 2.632% | 2.592% | 4.68% | 2.539% |  | 1.618% | 10.57% | 2.493% |  | 3.534% | 4.763% | 4.347% | 8.418% | 2.409% | 4.152% | 3.484%
| PDM3 |  | 31.43% |  | -14.29% | 49.6% | -40.02% | 53.79% | 4.723% | 0.1091% | 15.21% | -0.7298% | 23.04% |  | 16.84% | -58.89% | -10.56% |  | 30.03% | -3.526% | -32.05% | -7.441% | 0.1091% | -17.51% | 2.905%
| Z-Score (Absolute) |  | 1.23 | 4.3 | 0.57 | 1.94 | 1.58 | 2.11 | 0.18 | 0 | 0.59 | 0.03 | 0.9 | 2.28 | 0.66 | 2.32 | 0.42 | NA | 1.17 | 0.14 | 1.26 | 0.3 | 0 | 0.69 | 0.11
| Analytical results for Tb in OREAS 555b (Certified Value 0.81 ppm) | [28 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 20 | 23 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.5 | 0.75 | 1 | 1.1129 | 2 | 1.2723 | 0.84702 | NA | 0.9 | 0.62 | 0.68 | 0.83 | [15 empty] | 1
| 1 | 2 | 0.6 | 0.76 | 0.9 | 1.0916 | 1 | 1.2196 | 0.83519 | NA | 0.8 | 0.61 | 0.77 | 0.83 | [15 empty] | 6
| 1 | 3 | 0.5 | 0.78 | 1 | 1.0932 | 1 | 1.2521 | 0.82626 | NA | 0.9 | 0.61 | 0.69 | 0.84 | [15 empty] | 16
| 1 | 4 | 0.5 | 0.77 | 0.7 | 1.0811 | 1 | 1.2309 | 0.84859 | NA | 0.9 | 0.59 | 0.72 | 0.8 | [15 empty] | 0.81401
| 1 | 5 | 0.5 | 0.77 | 1 | 1.0471 | 1 | 1.2455 | 0.87432 | NA | 0.9 | 0.63 | 0.82 | 0.84 | [15 empty] | 134
| 1 | 6 | 0.5 | 0.77 | 1 | 1.1104 | 1 | 1.3817 | 0.87673 | NA | 0.9 | 0.65 | 0.75 | 0.78
| Mean |  | 0.51667 | 0.76667 | 0.93333 | 1.0894 | 1.1667 | 1.267 | 0.85135 |  | 0.88333 | 0.61833 | 0.73833 | 0.82
| Median |  | 0.5 | 0.77 | 1 | 1.0924 | 1 | 1.2488 | 0.84781 |  | 0.9 | 0.615 | 0.735 | 0.83
| Std Dev. |  | 0.040825 | 0.010328 | 0.12111 | 0.02398 | 0.40825 | 0.059003 | 0.020438 |  | 0.040825 | 0.020412 | 0.052694 | 0.024495
| Rel.Std.Dev. |  | 7.902% | 1.347% | 12.98% | 2.201% | 34.99% | 4.657% | 2.401% |  | 4.622% | 3.301% | 7.137% | 2.987%
| PDM3 |  | -36.53% | -5.816% | 14.66% | 33.83% | 43.32% | 55.65% | 4.588% |  | 8.516% | -24.04% | -9.297% | 0.736%
| Z-Score (Absolute) |  | NA | 0.44 | NA | 2.22 | NA | 3.69 | 0.26 | NA | NA | 1.67 | 0.67 | 0
Results from laboratories 2, 5 and 20 were removed due to their 0.1 ppm reading resolution. Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Te in OREAS 555b (Certified Value 0.06 ppm) | [28 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.06 | < 5 | < 0.2 | < 0.1 | 0.12 | 0.05919 | 0.05 | 0.07 | < 0.05 | 0.06 | 0.07 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | < 0.05 | < 0.2 | 0.08 | [7 empty] | 1
| 1 | 2 | 0.1 | < 5 | < 0.2 | < 0.1 | 0.13 | 0.07532 | 0.06 | 0.06 | 0.06 | < 0.05 | 0.09 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | < 0.05 | < 0.2 | 0.1 | [7 empty] | 7
| 1 | 3 | 0.06 | < 5 | < 0.2 | < 0.1 | 0.02 | 0.07061 | < 0.05 | 0.06 | < 0.05 | 0.06 | 0.08 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | < 0.05 | < 0.2 | 0.08 | [7 empty] | 16
| 1 | 4 | 0.06 | < 5 | < 0.2 | < 0.1 | 0.06 | 0.06608 | 0.06 | 0.06 | < 0.05 | 0.06 | 0.09 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | < 0.05 | < 0.2 | 0.13 | [7 empty] | 0.059818
| 1 | 5 | 0.05 | < 5 | < 0.2 | < 0.1 | 0.02 | 0.06847 | 0.06 | 0.05 | 0.11 | 0.09 | 0.09 | < 10 | < 0.1 | < 0.1 | < 5 | NA | 0.1 | < 0.05 | < 0.2 | 0.07 | [7 empty] | 135
| 1 | 6 | 0.06 | < 5 | < 0.2 | < 0.1 | < 0.01 | 0.06577 | < 0.05 | 0.06 | < 0.05 | 0.11 | 0.08 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | < 0.05 | < 0.2 | 0.11
| Mean |  | 0.065 |  |  |  | 0.07 | 0.067573 | 0.0575 | 0.06 | 0.085 | 0.076 | 0.083333 | [5 empty] | 0.1 |  |  | 0.095
| Median |  | 0.06 |  |  |  | 0.06 | 0.067275 | 0.06 | 0.06 | 0.085 | 0.06 | 0.085 | [5 empty] | 0.1 |  |  | 0.09
| Std Dev. |  | 0.017607 |  |  |  | 0.052915 | 0.0054 | 0.005 | 0.0063246 | 0.035355 | 0.023022 | 0.008165 | [8 empty] | 0.022583
| Rel.Std.Dev. |  | 27.09% |  |  |  | 75.59% | 7.991% | 8.696% | 10.54% | 41.59% | 30.29% | 9.798% | [8 empty] | 23.77%
| PDM3 |  | 8.663% |  |  |  | 17.02% | 12.97% | -3.875% | 0.3046% | 42.1% | 27.05% | 39.31% | [5 empty] | 67.17% |  |  | 58.82%
| Z-Score (Absolute) |  | 0.05 | 89.51 | 1.24 | 0.6 | 0.26 | 0.05 | 0.72 | 0.23 | 0.78 | 0.05 | 0.63 | 181.47 | 0.6 | 0.6 | 89.51 | NA | NA | 1.52 | 1.24 | 1.06
| Analytical results for Th in OREAS 555b (Certified Value 16 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 14.6 | 15.51 | 17 | 16 | 15.793 | 16.85 | 15.95 | 15.65 | 16.75 | 17.25 | 14 | 16.5836 | 17.8756 | 15.8 | 16.0995 | 16.32 | 15.1 | 16.9 | 14.64 | 15.6 | 17.09 | 16.1 | 14.8 |  |  |  |  | 1
| 1 | 2 | 14.4 | 15.3 | 16.5 | 16.3 | 15.6891 | 16.85 | 14.65 | 16.25 | 16.4 | 16.85 | 15 | 16.6541 | 17.1826 | 15.3 | 15.6709 | 15.26 | 14.79 | 12.1 | 14.76 | 15.4 | 16.59 | 15.6 | 14.6 |  |  |  |  | 8
| 1 | 3 | 13.7 | 15.34 | 16.2 | 16.6 | 15.8637 | 17.2 | 15.75 | 16.45 | 17.35 | 17.25 | 15 | 16.4725 | 17.2587 | 16 | 15.704 | 15.83 | 15.3 | 16.6 | 14.64 | 15.6 | 16.78 | 16.5 | 14.8 |  |  |  |  | 16
| 1 | 4 | 14.5 | 15.29 | 11.9 | 16.6 | 15.9424 | 16.55 | 14.5 | 15.7 | 16.85 | 14.85 | 16 | 16.585 | 17.8766 | 15.7 | 16.348 | 16.6 | 14.55 | 15.3 | 14.28 | 15.3 | 16.96 | 16.2 | 15.6 |  |  |  |  | 15.977
| 1 | 5 | 14.1 | 15.45 | 16.7 | 16.6 | 15.8573 | 17.05 | 15.25 | 15.7 | 16.05 | 17.65 | 15 | 16.63 | 17.9487 | 15.6 | 16.2518 | 15.89 | 14.52 | 14.5 | 14.97 | 17.6 | 17.04 | 16.5 | 14.8 |  |  |  |  | 136
| 1 | 6 | 13.9 | 15.34 | 17 | 16.5 | 16.2181 | 17.05 | 14.85 | 15.45 | 16.95 | 15.05 | 15 | 16.4959 | 17.8954 | 15.6 | 16.7765 | 15.7 | 15.47 | 14.6 | 15.52 | 18 | 16.93 | 16.3 | 15.4
| Mean |  | 14.2 | 15.3717 | 15.8833 | 16.4333 | 15.8939 | 16.925 | 15.1583 | 15.8667 | 16.725 | 16.4833 | 15 | 16.5702 | 17.6729 | 15.6667 | 16.1418 | 15.9333 | 14.955 | 15 | 14.8017 | 16.25 | 16.8983 | 16.2 | 15
| Median |  | 14.25 | 15.34 | 16.6 | 16.55 | 15.8605 | 16.95 | 15.05 | 15.7 | 16.8 | 17.05 | 15 | 16.5843 | 17.8761 | 15.65 | 16.1756 | 15.86 | 14.945 | 14.95 | 14.7 | 15.6 | 16.945 | 16.25 | 14.8
| Std Dev. |  | 0.35777 | 0.088412 | 1.9753 | 0.24221 | 0.17989 | 0.22749 | 0.59533 | 0.39073 | 0.45139 | 1.216 | 0.63246 | 0.07223 | 0.35218 | 0.23381 | 0.41776 | 0.47209 | 0.39642 | 1.7367 | 0.41725 | 1.2128 | 0.18476 | 0.33466 | 0.4
| Rel.Std.Dev. |  | 2.52% | 0.5752% | 12.44% | 1.474% | 1.132% | 1.344% | 3.927% | 2.463% | 2.699% | 7.377% | 4.216% | 0.4359% | 1.993% | 1.492% | 2.588% | 2.963% | 2.651% | 11.58% | 2.819% | 7.464% | 1.093% | 2.066% | 2.667%
| PDM3 |  | -11.12% | -3.789% | -0.5862% | 2.856% | -0.52% | 5.934% | -5.124% | -0.6905% | 4.682% | 3.169% | -6.115% | 3.713% | 10.61% | -1.942% | 1.031% | -0.2732% | -6.397% | -6.115% | -7.356% | 1.709% | 5.767% | 1.396% | -6.115%
| Z-Score (Absolute) |  | 1.88 | 0.6 | 0.03 | 0.57 | 0.02 | 1.11 | 0.83 | 0.05 | 0.89 | 0.63 | NA | 0.72 | 1.93 | 0.27 | 0.25 | 0.02 | 1.05 | 1 | 1.22 | 0.37 | 1.08 | 0.31 | 1
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 555b (Certified Value 0.274 wt.%) | [28 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.4 | 0.308 | 0.12 | 0.2391 | 0.261 | 0.18 | 0.45212 | 0.304 | 0.288 | 0.304 | 0.285 | 0.357 | 0.42 | 0.37758 | 0.21594 | 0.141 | 0.28936 | NA | 0.34 | 0.13 | 0.25 | 0.14 | 0.235 | 0.2524 | 0.236 | 0.34 |  | 1
| 1 | 2 | 0.4 | 0.295 | 0.12 | 0.242 | 0.259 | 0.16 | 0.45447 | 0.306 | 0.294 | 0.323 | 0.283 | 0.355 | 0.42 | 0.36341 | 0.21739 | 0.143 | 0.29131 | NA | 0.33 | 0.12 | 0.237 | 0.137 | 0.237 | 0.2613 | 0.233 | 0.34 |  | 25
| 1 | 3 | 0.39 | 0.316 | 0.12 | 0.2324 | 0.259 | 0.16 | 0.45537 | 0.314 | 0.297 | 0.32 | 0.287 | 0.361 | 0.426 | 0.37756 | 0.2101 | 0.15 | 0.28625 | NA | 0.33 | 0.13 | 0.252 | 0.149 | 0.241 | 0.2695 | 0.23 | 0.32 |  | 16
| 1 | 4 | 0.4 | 0.326 | 0.12 | 0.2062 | 0.257 | 0.17 | 0.46733 | 0.309 | 0.295 | 0.323 | 0.29 | 0.359 | 0.432 | 0.37991 | 0.21809 | 0.146 | 0.29134 | NA | 0.33 | 0.12 | 0.236 | 0.124 | 0.288 | 0.2682 | 0.223 | 0.34 |  | 0.27375
| 1 | 5 | 0.39 | 0.332 | 0.12 | 0.2466 | 0.251 | 0.18 | 0.46406 | 0.308 | 0.289 | 0.323 | 0.268 | 0.344 | 0.426 | 0.35569 | 0.21298 | 0.146 | 0.29296 | NA | 0.33 | 0.12 | 0.241 | 0.137 | 0.264 | 0.2678 | 0.234 | 0.35 |  | 137
| 1 | 6 | 0.39 | 0.324 | 0.13 | 0.2118 | 0.268 | 0.16 | 0.45248 | 0.31 | 0.285 | 0.313 | 0.298 | 0.359 | 0.42 | 0.36442 | 0.21527 | 0.143 | 0.29264 | NA | 0.32 | 0.13 | 0.235 | 0.136 | 0.257 | 0.2507 | 0.227 | 0.35
| Mean |  | 0.395 | 0.31683 | 0.12167 | 0.22968 | 0.25917 | 0.16833 | 0.45764 | 0.3085 | 0.29133 | 0.31767 | 0.28517 | 0.35583 | 0.424 | 0.36976 | 0.21496 | 0.14483 | 0.29064 |  | 0.33 | 0.125 | 0.24183 | 0.13717 | 0.25367 | 0.26165 | 0.2305 | 0.34
| Median |  | 0.395 | 0.32 | 0.12 | 0.23575 | 0.259 | 0.165 | 0.45492 | 0.3085 | 0.2915 | 0.3215 | 0.286 | 0.358 | 0.423 | 0.37099 | 0.2156 | 0.1445 | 0.29132 |  | 0.33 | 0.125 | 0.239 | 0.137 | 0.249 | 0.26455 | 0.2315 | 0.34
| Std Dev. |  | 0.0054772 | 0.013571 | 0.0040825 | 0.016764 | 0.0055287 | 0.0098319 | 0.0064413 | 0.0034496 | 0.0046762 | 0.0077374 | 0.0099079 | 0.0061455 | 0.004899 | 0.0099176 | 0.0029765 | 0.0031885 | 0.0024994 |  | 0.0063246 | 0.0054772 | 0.007414 | 0.0080353 | 0.020412 | 0.0083419 | 0.0048477 | 0.010954
| Rel.Std.Dev. |  | 1.387% | 4.283% | 3.355% | 7.299% | 2.133% | 5.841% | 1.408% | 1.118% | 1.605% | 2.436% | 3.474% | 1.727% | 1.155% | 2.682% | 1.385% | 2.202% | 0.86% |  | 1.917% | 4.382% | 3.066% | 5.858% | 8.047% | 3.188% | 2.103% | 3.222%
| PDM3 |  | 44.29% | 15.74% | -55.56% | -16.1% | -5.328% | -38.51% | 67.17% | 12.69% | 6.423% | 16.04% | 4.17% | 29.98% | 54.89% | 35.07% | -21.48% | -47.09% | 6.17% |  | 20.55% | -54.34% | -11.66% | -49.89% | -7.337% | -4.42% | -15.8% | 24.2%
| Z-Score (Absolute) |  | 1.35 | 0.39 | 2.01 | 0.68 | 0.32 | 1.44 | 2.12 | 0.29 | 0.08 | 0.4 | 0 | 0.87 | 1.7 | 1.04 | 0.86 | 1.73 | 0.07 | NA | 0.55 | 1.97 | 0.53 | 1.82 | 0.39 | 0.29 | 0.67 | 0.67
| Analytical results for Tl in OREAS 555b (Certified Value 0.99 ppm) | [28 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 0.89 | 1 | 1.02 | 1.03 | 1.0191 | 0.9 | 0.9 | 0.92 | 1.02 | 1.04 | 1.0069 | 1.02 | 1.0116 | 0.365 | 0.98 | 2 | 0.9 | 0.9 | 1.05 | 1.11 | 0.92 | [6 empty] | 1
| 1 | 2 | 0.94 | 1.03 | 1 | 1.04 | 1.0129 | 0.91 | 0.89 | 1 | 1.04 | 0.97 | 1.0084 | 1.01 | 0.99095 | 0.328 | 0.97 | < 2 | 0.9 | 0.89 | 1.05 | 1.03 | 0.95 | [6 empty] | 26
| 1 | 3 | 0.85 | 1.02 | 1.01 | 1.06 | 1.0337 | 0.92 | 0.92 | 1.1 | 1.03 | 1 | 1.0203 | 1.06 | 0.96898 | 0.337 | 0.99 | < 2 | 0.8 | 0.9 | 1.02 | 1.07 | 0.92 | [6 empty] | 16
| 1 | 4 | 0.9 | 1.05 | 0.74 | 1.05 | 0.99438 | 0.93 | 0.87 | 0.99 | 1.02 | 0.94 | 1.0267 | 1.03 | 0.99954 | 0.356 | 0.99 | 2 | 0.9 | 0.87 | 1.03 | 1.09 | 0.97 | [6 empty] | 0.9867
| 1 | 5 | 0.88 | 1.04 | 0.99 | 1.06 | 1.0027 | 0.92 | 0.92 | 0.97 | 1.05 | 1.04 | 1.033 | 1.05 | 0.98448 | 0.311 | 0.98 | < 2 | 0.9 | 0.88 | 1.14 | 1.1 | 0.94 | [6 empty] | 138
| 1 | 6 | 0.84 | 1.03 | 0.99 | 1.05 | 1.0179 | 0.92 | 0.89 | 1.04 | 1.03 | 0.93 | 1.0093 | 1.04 | 1.041 | 0.243 | 0.98 | 2 | 1.1 | 0.93 | 1.15 | 1.06 | 0.94
| Mean |  | 0.88333 | 1.0283 | 0.95833 | 1.0483 | 1.0134 | 0.91667 | 0.89833 | 1.0033 | 1.0317 | 0.98667 | 1.0174 | 1.035 | 0.99941 | 0.32333 | 0.98167 | 2 | 0.91667 | 0.895 | 1.0733 | 1.0767 | 0.94
| Median |  | 0.885 | 1.03 | 0.995 | 1.05 | 1.0154 | 0.92 | 0.895 | 0.995 | 1.03 | 0.985 | 1.0148 | 1.035 | 0.99525 | 0.3325 | 0.98 | 2 | 0.9 | 0.895 | 1.05 | 1.08 | 0.94
| Std Dev. |  | 0.036148 | 0.017224 | 0.10759 | 0.01169 | 0.013727 | 0.010328 | 0.019408 | 0.061536 | 0.01169 | 0.048028 | 0.010905 | 0.018708 | 0.024877 | 0.043857 | 0.0075277 | 0 | 0.098319 | 0.020736 | 0.056804 | 0.029439 | 0.018974
| Rel.Std.Dev. |  | 4.092% | 1.675% | 11.23% | 1.115% | 1.354% | 1.127% | 2.16% | 6.133% | 1.133% | 4.868% | 1.072% | 1.808% | 2.489% | 13.56% | 0.7668% | 0% | 10.73% | 2.317% | 5.292% | 2.734% | 2.018%
| PDM3 |  | -10.48% | 4.22% | -2.875% | 6.246% | 2.711% | -7.098% | -8.956% | 1.686% | 4.557% | -0.003299% | 3.114% | 4.895% | 1.288% | -67.23% | -0.51% | 102.7% | -7.098% | -9.294% | 8.78% | 9.118% | -4.733%
| Z-Score (Absolute) |  | 1.56 | 0.5 | 0.49 | 0.78 | 0.29 | 1.08 | 1.34 | 0.15 | 0.55 | 0.09 | 0.35 | 0.6 | 0.09 | 9.51 | 0.16 | NA | 1.08 | 1.39 | 1.14 | 1.19 | 0.75
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Tm in OREAS 555b (Certified Value 0.42 ppm) | [28 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 16 | 18 | 19 | 23 | 25 | 26 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [17 empty] | 2
2
| 1 | 1 | 0.3 | 0.39 | 0.5 | 0.53596 | 0.48428 | 0.43262 | 0.35 | 0.35 | 0.35 | 0.45 | [17 empty] | 1
| 1 | 2 | 0.3 | 0.39 | 0.5 | 0.53276 | 0.45801 | 0.43693 | 0.32 | 0.35 | 0.41 | 0.45 | [17 empty] | 10
| 1 | 3 | 0.3 | 0.4 | 0.4 | 0.52374 | 0.45267 | 0.42868 | 0.34 | 0.34 | 0.36 | 0.44 | [17 empty] | 16
| 1 | 4 | 0.3 | 0.41 | 0.4 | 0.51353 | 0.47625 | 0.43917 | 0.36 | 0.32 | 0.38 | 0.44 | [17 empty] | 0.41895
| 1 | 5 | 0.3 | 0.43 | 0.5 | 0.51113 | 0.43855 | 0.4457 | 0.34 | 0.35 | 0.42 | 0.46 | [17 empty] | 139
| 1 | 6 | 0.3 | 0.4 | 0.5 | 0.53848 | 0.46604 | 0.44531 | 0.33 | 0.36 | 0.39 | 0.44
| Mean |  | 0.3 | 0.40333 | 0.46667 | 0.52593 | 0.46263 | 0.43807 | 0.34 | 0.345 | 0.385 | 0.44667
| Median |  | 0.3 | 0.4 | 0.5 | 0.52825 | 0.46202 | 0.43805 | 0.34 | 0.35 | 0.385 | 0.445
| Std Dev. |  | 0 | 0.015055 | 0.05164 | 0.011682 | 0.016531 | 0.0067977 | 0.014142 | 0.013784 | 0.027386 | 0.008165
| Rel.Std.Dev. |  | 0% | 3.733% | 11.07% | 2.221% | 3.573% | 1.552% | 4.159% | 3.995% | 7.113% | 1.828%
| PDM3 |  | -28.39% | -3.729% | 11.39% | 25.53% | 10.43% | 4.562% | -18.85% | -17.65% | -8.105% | 6.615%
| Z-Score (Absolute) |  | NA | 0.3 | NA | 1.83 | 0.73 | 0.3 | 1.4 | 1.32 | 0.62 | 0.45
Results from laboratories 2 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 555b (Certified Value 9.24 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 7.9 | 9.31 | 9.8 | 9.58 | 9.1384 | 9.8 | 8.5 | 8.5 | 9.3 | 9.4 | < 2 | 9.4662 | 10.4676 | 9.65 | 8.5174 | 8.32 | 9.31 | 9.6 | 8.83 | 10.3 | 9.72 | 9.6 | 8.5 |  |  |  |  | 1
| 1 | 2 | 8.6 | 9.38 | 9.7 | 9.67 | 9.1831 | 10.1 | 8 | 9.1 | 9.1 | 9.5 | 2 | 9.4504 | 10.8939 | 9.39 | 8.3258 | 7.64 | 9.2 | 9.4 | 8.78 | 10 | 9.5 | 9.1 | 8.6 |  |  |  |  | 11
| 1 | 3 | 7.6 | 9.4 | 9.6 | 9.84 | 9.206 | 9.9 | 8.7 | 9.1 | 9.7 | 9.4 | 5 | 9.466 | 10.1633 | 9.23 | 8.3486 | 7.96 | 9.5 | 9.2 | 9 | 10 | 9.64 | 9.9 | 8.3 |  |  |  |  | 16
| 1 | 4 | 7.9 | 9.26 | 7 | 9.8 | 9.1825 | 9.6 | 7.9 | 8.7 | 9.3 | 8.9 | < 2 | 9.3389 | 10.8112 | 9.49 | 8.3898 | 8.41 | 9.45 | 9.3 | 8.57 | 10.1 | 9.63 | 9.5 | 8.8 |  |  |  |  | 9.2368
| 1 | 5 | 7.7 | 9.43 | 9.5 | 9.8 | 9.2288 | 10.3 | 8.5 | 8.8 | 8.9 | 9.6 | 2 | 9.3289 | 10.9314 | 9.51 | 8.4561 | 8.01 | 9.31 | 9 | 8.89 | 11.4 | 9.65 | 9.8 | 8.5 |  |  |  |  | 140
| 1 | 6 | 7.5 | 9.36 | 10.1 | 9.88 | 9.2827 | 10.1 | 8.3 | 8.6 | 9.4 | 9.2 | 5 | 9.3689 | 10.0021 | 9.5 | 8.6583 | 7.87 | 9.47 | 9.4 | 9.01 | 11.7 | 9.7 | 9.6 | 8.7
| Mean |  | 7.8667 | 9.3567 | 9.2833 | 9.7617 | 9.2036 | 9.9667 | 8.3167 | 8.8 | 9.2833 | 9.3333 | 3.5 | 9.4032 | 10.5449 | 9.4617 | 8.4493 | 8.035 | 9.3733 | 9.3167 | 8.8467 | 10.5833 | 9.64 | 9.5833 | 8.5667
| Median |  | 7.8 | 9.37 | 9.65 | 9.8 | 9.1946 | 10 | 8.4 | 8.75 | 9.3 | 9.4 | 3.5 | 9.4096 | 10.6394 | 9.495 | 8.423 | 7.985 | 9.38 | 9.35 | 8.86 | 10.2 | 9.645 | 9.6 | 8.55
| Std Dev. |  | 0.39328 | 0.062183 | 1.1374 | 0.11356 | 0.049011 | 0.25033 | 0.31252 | 0.25298 | 0.27142 | 0.25033 | 1.7321 | 0.064777 | 0.39707 | 0.14063 | 0.12434 | 0.28683 | 0.11776 | 0.20412 | 0.1633 | 0.76267 | 0.077201 | 0.27869 | 0.17512
| Rel.Std.Dev. |  | 4.999% | 0.6646% | 12.25% | 1.163% | 0.5325% | 2.512% | 3.758% | 2.875% | 2.924% | 2.682% | 49.49% | 0.6889% | 3.766% | 1.486% | 1.472% | 3.57% | 1.256% | 2.191% | 1.846% | 7.206% | 0.8008% | 2.908% | 2.044%
| PDM3 |  | -14.83% | 1.298% | 0.5039% | 5.682% | -0.3594% | 7.902% | -9.962% | -4.729% | 0.5039% | 1.045% | -62.11% | 1.802% | 14.16% | 2.435% | -8.525% | -13.01% | 1.478% | 0.8648% | -4.224% | 14.58% | 4.365% | 3.752% | -7.255%
| Z-Score (Absolute) |  | 2.62 | 0.06 | 0.07 | 0.78 | 0.22 | 1.15 | 1.81 | 0.94 | 0.07 | 0.01 | NA | 0.14 | 2.19 | 0.25 | 1.57 | 2.31 | 0.09 | 0.01 | 0.86 | 2.26 | 0.57 | 0.46 | 1.36
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 555b (Certified Value 357 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 352.3 | 376 | 352 | 347 | 393 | 324 | 348.7047 | 344 | 330 | 330 | 325 | 367 | 333 | 360 | 381.8701 | 375.08 | 355 | 350.6022 | 380 | 360 | 322 | 357 | 344 | 335 | 362 | 356 | 364 | 1
| 1 | 2 | 348.1 | 359 | 355 | 355 | 387 | 337 | 352.7433 | 338 | 338 | 356 | 336 | 345 | 333 | 353 | 374.9054 | 371.76 | 356 | 356.2733 | 384 | 367 | 316 | 351 | 339 | 341 | 361 | 351 | 366 | 38
| 1 | 3 | 344.9 | 368 | 359 | 347 | 388 | 337 | 359.3444 | 349 | 334 | 352 | 348 | 359 | 335 | 368 | 381.093 | 375.92 | 353 | 346.8622 | 388 | 382 | 322 | 351 | 327 | 356 | 344 | 353 | 356 | 16
| 1 | 4 | 352.2 | 381 | 366 | 345 | 286 | 350 | 357.2237 | 344 | 336 | 362 | 339 | 360 | 332 | 366 | 378.1303 | 371.13 | 343 | 359.0692 | 381 | 377 | 315 | 345 | 330 | 371 | 341 | 351 | 368 | 356.6079
| 1 | 5 | 341.9 | 387 | 360 | 364 | 375 | 358 | 356.3987 | 342 | 343 | 358 | 334 | 349 | 334 | 360 | 379.9578 | 374.14 | 355 | 359.1923 | 380 | 384 | 320 | 356 | 331 | 357 | 355 | 357 | 369 | 141
| 1 | 6 | 344.4 | 373 | 361 | 364 | 378 | 368 | 349.4695 | 341 | 323 | 354 | 348 | 350 | 336 | 361 | 373.0574 | 379.42 | 351 | 356.0917 | 380 | 385 | 314 | 361 | 332 | 370 | 356 | 353 | 360
| Mean |  | 347.3 | 374 | 358.8333 | 353.6667 | 367.8333 | 345.6667 | 353.9807 | 343 | 334 | 352 | 338.3333 | 355 | 333.8333 | 361.3333 | 378.169 | 374.575 | 352.1667 | 354.6818 | 382.1667 | 375.8333 | 318.1667 | 353.5 | 333.8333 | 355 | 353.1667 | 353.5 | 363.8333
| Median |  | 346.5 | 374.5 | 359.5 | 351 | 382.5 | 343.5 | 354.571 | 343 | 335 | 355 | 337.5 | 354.5 | 333.5 | 360.5 | 379.0441 | 374.61 | 354 | 356.1825 | 380.5 | 379.5 | 318 | 353.5 | 331.5 | 356.5 | 355.5 | 353 | 365
| Std Dev. |  | 4.3123 | 9.8387 | 4.8751 | 8.7101 | 40.6419 | 16.0582 | 4.3555 | 3.6878 | 6.8993 | 11.3137 | 8.8242 | 8.3187 | 1.472 | 5.2789 | 3.5279 | 3.0173 | 4.8339 | 4.9377 | 3.2506 | 10.1866 | 3.6009 | 5.648 | 6.3692 | 14.7105 | 8.7502 | 2.51 | 4.9967
| Rel.Std.Dev. |  | 1.242% | 2.631% | 1.359% | 2.463% | 11.05% | 4.646% | 1.23% | 1.075% | 2.066% | 3.214% | 2.608% | 2.343% | 0.4409% | 1.461% | 0.9329% | 0.8055% | 1.373% | 1.392% | 0.8506% | 2.71% | 1.132% | 1.598% | 1.908% | 4.144% | 2.478% | 0.71% | 1.373%
| PDM3 |  | -2.61% | 4.877% | 0.624% | -0.8248% | 3.148% | -3.068% | -0.7367% | -3.816% | -6.34% | -1.292% | -5.125% | -0.4509% | -6.386% | 1.325% | 6.046% | 5.038% | -1.245% | -0.5401% | 7.167% | 5.391% | -10.78% | -0.8715% | -6.386% | -0.4509% | -0.965% | -0.8715% | 2.026%
| Z-Score (Absolute) |  | 0.54 | 1.71 | 0.44 | 0 | 1.19 | 0.67 | 0.03 | 0.9 | 1.66 | 0.14 | 1.29 | 0.11 | 1.67 | 0.65 | 2.07 | 1.76 | 0.13 | 0.09 | 2.4 | 1.87 | 2.99 | 0.01 | 1.67 | 0.11 | 0.04 | 0.01 | 0.86
| Analytical results for W in OREAS 555b (Certified Value 2.78 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 5 | 2.8 | < 10 | 2.6 | 3.4 | 2.4 | 3.1281 | 2.7 | 2.5 | 2.9 | 2.7 | 3.1 | 12 | 2.8308 | 2.1 | 2.75 | NA | 2.8 | < 10 | 2.8 | 2.5 | 2.9 | 2.7 | 3.1 | 2.9 |  |  | 1
| 1 | 2 | < 5 | 3 | < 10 | 2.6 | 3.2 | 2.4 | 3.2243 | 2.8 | 2.6 | 2.9 | 2.7 | 2.9 | 23 | 2.8083 | 2.4 | 2.7266 | NA | 2.8 | < 10 | 2.6 | 2.4 | 3.2 | 2.8 | 2.5 | 2.9 |  |  | 12
| 1 | 3 | < 5 | 2.7 | < 10 | 2.7 | 3.2 | 2.4 | 3.1281 | 2.9 | 2.6 | 2.9 | 2.8 | 3.1 | < 1 | 2.8267 | 2.5 | 2.7956 | NA | 2.9 | < 10 | 2.6 | 2.4 | 2.9 | 2.9 | 2.5 | 2.9 |  |  | 16
| 1 | 4 | < 5 | 2.8 | < 10 | 2.5 | 2.3 | 2.4 | 3.092 | 2.7 | 2.5 | 3.1 | 2.7 | 2.8 | 18 | 2.8593 | 2.3 | 2.7057 | NA | 2.8 | < 10 | 2.7 | 2.3 | 3.1 | 3 | 2.5 | 3 |  |  | 2.7793
| 1 | 5 | < 5 | 2.8 | < 10 | 2.6 | 3.2 | 2.3 | 3.1717 | 2.8 | 2.7 | 2.9 | 2.6 | 3.2 | 14 | 2.8537 | 2.4 | 2.685 | NA | 2.8 | < 10 | 2.7 | 2.3 | 3.4 | 2.9 | 2.7 | 2.9 |  |  | 142
| 1 | 6 | < 5 | 2.7 | < 10 | 3 | 3.3 | 2.4 | 3.1934 | 2.7 | 2.6 | 2.8 | 2.8 | 3 | 15 | 2.7922 | 2.4 | 2.7875 | NA | 2.8 | < 10 | 2.6 | 2.4 | 3.2 | 2.8 | 2.7 | 2.9
| Mean |  |  | 2.8 |  | 2.6667 | 3.1 | 2.3833 | 3.1563 | 2.7667 | 2.5833 | 2.9167 | 2.7167 | 3.0167 | 16.4 | 2.8285 | 2.35 | 2.7417 |  | 2.8167 |  | 2.6667 | 2.3833 | 3.1167 | 2.85 | 2.6667 | 2.9167
| Median |  |  | 2.8 |  | 2.6 | 3.2 | 2.4 | 3.1499 | 2.75 | 2.6 | 2.9 | 2.7 | 3.05 | 15 | 2.8288 | 2.4 | 2.7383 |  | 2.8 |  | 2.65 | 2.4 | 3.15 | 2.85 | 2.6 | 2.9
| Std Dev. |  |  | 0.10954 |  | 0.17512 | 0.4 | 0.040825 | 0.048925 | 0.08165 | 0.075277 | 0.098319 | 0.075277 | 0.1472 | 4.2778 | 0.025729 | 0.13784 | 0.044313 |  | 0.040825 |  | 0.08165 | 0.075277 | 0.19408 | 0.10488 | 0.23381 | 0.040825
| Rel.Std.Dev. |  |  | 3.912% |  | 6.567% | 12.9% | 1.713% | 1.55% | 2.951% | 2.914% | 3.371% | 2.771% | 4.879% | 26.08% | 0.9096% | 5.866% | 1.616% |  | 1.449% |  | 3.062% | 3.158% | 6.227% | 3.68% | 8.768% | 1.4%
| PDM3 |  |  | 0.7439% |  | -4.053% | 11.54% | -14.25% | 13.56% | -0.4555% | -7.052% | 4.942% | -2.254% | 8.54% | 490.1% | 1.769% | -15.45% | -1.353% |  | 1.344% |  | -4.053% | -14.25% | 12.14% | 2.543% | -4.053% | 4.942%
| Z-Score (Absolute) |  | 1.52 | 0 | 11.13 | 0.67 | 1.52 | 2.11 | 1.8 | 0.17 | 1.1 | 0.59 | 0.42 | 1.1 | NA | 0.14 | 2.28 | 0.29 | NA | 0.08 | 11.13 | 0.67 | 2.11 | 1.6 | 0.25 | 0.67 | 0.59
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 555b (Certified Value 24.4 ppm) | [28 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 24.1 | 18.6 | 23.91 | 29.7 | 24.2 | 34.8051 | 23.3 | 23.4 | 23.2 | 24.5 | 25.1 | 19.49 | 26 | 25.9768 | 24.8 | 29.884 | 20.84 | 25.2 | 21.5 | 26.3 | 19.7 | 20.5 | 26.07 | 24.3 | 23.3 |  |  | 1
| 1 | 2 | 25.3 | 19.4 | 25.95 | 27.9 | 24.5 | 34.0278 | 24.2 | 22.3 | 24.9 | 24.2 | 25 | 19.27 | 25 | 25.3344 | 24.1 | 24.0774 | 21.02 | 24.4 | 20.1 | 22.6 | 19.9 | 23.8 | 26.11 | 23.9 | 23.6 |  |  | 13
| 1 | 3 | 24.1 | 19.2 | 24.15 | 30.4 | 24.3 | 34.39 | 24.5 | 23.7 | 24.4 | 24.6 | 25.1 | 19.35 | 25 | 25.4915 | 25.4 | 23.6205 | 21.32 | 25.4 | 21.4 | 26.8 | 18.2 | 21 | 26.31 | 25.1 | 22.5 |  |  | 16
| 1 | 4 | 24.7 | 19.9 | 24.33 | 22.2 | 25.1 | 33.9263 | 23.5 | 21.7 | 25.5 | 23.9 | 23.6 | 19.38 | 26 | 25.9417 | 23.9 | 24.12 | 21 | 25 | 21.1 | 22.1 | 18.4 | 22.2 | 25.07 | 25.3 | 23.5 |  |  | 24.4386
| 1 | 5 | 24.8 | 19.1 | 24.51 | 29.5 | 25.5 | 33.9002 | 24.5 | 24.1 | 23.7 | 23.6 | 24.8 | 19.29 | 25 | 25.8903 | 25.4 | 24.1369 | 20.95 | 25.1 | 20.5 | 22.5 | 18.9 | 24.5 | 25.5 | 24.5 | 23.4 |  |  | 143
| 1 | 6 | 24.5 | 18.9 | 24.88 | 30.1 | 25.8 | 34.3886 | 23.9 | 22.8 | 24.4 | 24.1 | 24.5 | 19.32 | 24 | 25.6406 | 24.8 | 23.9919 | 20.6 | 25.3 | 21.5 | 23.6 | 19.6 | 22.5 | 24.94 | 24.4 | 23.8
| Mean |  | 24.5833 | 19.1833 | 24.6217 | 28.3 | 24.9 | 34.2397 | 23.9833 | 23 | 24.35 | 24.15 | 24.6833 | 19.35 | 25.1667 | 25.7126 | 24.7333 | 24.9718 | 20.955 | 25.0667 | 21.0167 | 23.9833 | 19.1167 | 22.4167 | 25.6667 | 24.5833 | 23.35
| Median |  | 24.6 | 19.15 | 24.42 | 29.6 | 24.8 | 34.2082 | 24.05 | 23.1 | 24.4 | 24.15 | 24.9 | 19.335 | 25 | 25.7655 | 24.8 | 24.0987 | 20.975 | 25.15 | 21.25 | 23.1 | 19.25 | 22.35 | 25.785 | 24.45 | 23.45
| Std Dev. |  | 0.45789 | 0.4446 | 0.72907 | 3.1119 | 0.66633 | 0.35293 | 0.50761 | 0.90333 | 0.82158 | 0.37283 | 0.57764 | 0.079246 | 0.75277 | 0.26495 | 0.6314 | 2.4141 | 0.23611 | 0.3559 | 0.58793 | 2.0547 | 0.71949 | 1.5484 | 0.58051 | 0.52313 | 0.45056
| Rel.Std.Dev. |  | 1.863% | 2.318% | 2.961% | 11% | 2.676% | 1.031% | 2.117% | 3.928% | 3.374% | 1.544% | 2.34% | 0.4095% | 2.991% | 1.03% | 2.553% | 9.667% | 1.127% | 1.42% | 2.797% | 8.567% | 3.764% | 6.908% | 2.262% | 2.128% | 1.93%
| PDM3 |  | 0.5923% | -21.5% | 0.7491% | 15.8% | 1.888% | 40.1% | -1.863% | -5.887% | -0.3625% | -1.181% | 1.001% | -20.82% | 2.979% | 5.213% | 1.206% | 2.182% | -14.25% | 2.57% | -14% | -1.863% | -21.78% | -8.273% | 5.025% | 0.5923% | -4.454%
| Z-Score (Absolute) |  | 0 | 6.07 | 0.04 | 4.18 | 0.36 | 10.85 | 0.67 | 1.78 | 0.26 | 0.49 | 0.11 | 5.88 | 0.66 | 1.27 | 0.17 | 0.44 | 4.08 | 0.54 | 4.01 | 0.67 | 6.14 | 2.44 | 1.22 | 0 | 1.39
| Analytical results for Yb in OREAS 555b (Certified Value 2.88 ppm) | [28 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 16 | 18 | 19 | 20 | 23 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 2.6 | 2.84 | 3.2 | 3.3001 | 3.2 | 3.3367 | 2.9337 | 2.66 | 2.9 | 2.3 | 2.39 | 3.13 | [15 empty] | 1
| 1 | 2 | 2.6 | 2.93 | 3 | 3.2393 | 3.1 | 3.149 | 2.8224 | 2.68 | 3.1 | 2.3 | 2.85 | 3.1 | [15 empty] | 14
| 1 | 3 | 2.3 | 2.86 | 3.2 | 3.3116 | 3.2 | 3.1953 | 2.7748 | 2.71 | 3 | 2.3 | 2.43 | 3.21 | [15 empty] | 16
| 1 | 4 | 2.5 | 2.91 | 2.4 | 3.1564 | 3.2 | 2.9765 | 2.8697 | 2.66 | 2.8 | 2.3 | 2.63 | 3.09 | [15 empty] | 2.8839
| 1 | 5 | 2.5 | 2.87 | 3.2 | 3.1556 | 3.1 | 3.1385 | 2.9292 | 2.66 | 2.8 | 2.2 | 2.8 | 3.09 | [15 empty] | 144
| 1 | 6 | 2.4 | 2.93 | 3.3 | 3.3612 | 3.3 | 3.1458 | 2.9339 | 2.59 | 2.9 | 2.3 | 2.69 | 3.04
| Mean |  | 2.4833 | 2.89 | 3.05 | 3.254 | 3.1833 | 3.157 | 2.8773 | 2.66 | 2.9167 | 2.2833 | 2.6317 | 3.11
| Median |  | 2.5 | 2.89 | 3.2 | 3.2697 | 3.2 | 3.1474 | 2.8994 | 2.66 | 2.9 | 2.3 | 2.66 | 3.095
| Std Dev. |  | 0.1169 | 0.038471 | 0.33317 | 0.085287 | 0.075277 | 0.11564 | 0.06732 | 0.039497 | 0.1169 | 0.040825 | 0.18894 | 0.056921
| Rel.Std.Dev. |  | 4.708% | 1.331% | 10.92% | 2.621% | 2.365% | 3.663% | 2.34% | 1.485% | 4.008% | 1.788% | 7.179% | 1.83%
| PDM3 |  | -13.89% | 0.2122% | 5.76% | 12.83% | 10.38% | 9.469% | -0.2292% | -7.763% | 1.137% | -20.82% | -8.746% | 7.841%
| Z-Score (Absolute) |  | 1.14 | 0.04 | 0.4 | 0.95 | 0.76 | 0.69 | 0.07 | 0.66 | 0.04 | 1.68 | 0.74 | 0.56
| Analytical results for Zn in OREAS 555b (Certified Value 29 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 35.61 | 30.6 | 29.9 | 30 | 30 | 25.9 | 30.2571 | 28 | 27 | 62 | 27 | 31 | 28.24 | 129 | 28.4758 | 28 | 28.0279 | 30 | 31.4 | 32.9 | 26 | 25.9 | 27 | 30 | 26 |  |  | 1
| 1 | 2 | 29.35 | 35.1 | 30 | 30 | 32 | 26.9 | 28.5687 | 29 | 27 | 30 | 27 | 30 | 25.97 | 121 | 29.2741 | 29 | 27.9595 | 30 | 30 | 33.2 | 25 | 30.2 | 26 | 29 | 27 |  |  | 15
| 1 | 3 | 33.28 | 30.5 | 28.6 | 30 | 30 | 27.3 | 28.4262 | 28 | 28 | 30 | 28 | 29 | 26.48 | 125 | 28.3323 | 28 | 28.0295 | 31 | 30.5 | 32.6 | 26 | 25.8 | 26 | 31 | 26 |  |  | 16
| 1 | 4 | 33.56 | 31 | 30.4 | 29 | 30 | 26.1 | 30.0334 | 28 | 28 | 30 | 27 | 29 | 26.72 | 127 | 29.23 | 27 | 28.0837 | 31 | 30.5 | 32.5 | 26 | 26.2 | 25 | 29 | 29 |  |  | 29.0115
| 1 | 5 | 30.86 | 31.5 | 30.1 | 31 | 30 | 28 | 31.5414 | 27 | 28 | 32 | 26 | 30 | 26.35 | 121 | 29.209 | 29 | 28.6109 | 31 | 31.1 | 33.3 | 25 | 28.5 | 27 | 30 | 26 |  |  | 145
| 1 | 6 | 41.47 | 31.1 | 28.8 | 31 | 30 | 28.9 | 30.6744 | 30 | 28 | 30 | 28 | 30 | 26.84 | 118 | 29.324 | 28 | 28.03 | 31 | 30.9 | 33.8 | 26 | 27.3 | 28 | 29 | 27
| Mean |  | 34.0217 | 31.6333 | 29.6333 | 30.1667 | 30.3333 | 27.1833 | 29.9169 | 28.3333 | 27.6667 | 35.6667 | 27.1667 | 29.8333 | 26.7667 | 123.5 | 28.9742 | 28.1667 | 28.1236 | 30.6667 | 30.7333 | 33.05 | 25.6667 | 27.3167 | 26.5 | 29.6667 | 26.8333
| Median |  | 33.42 | 31.05 | 29.95 | 30 | 30 | 27.1 | 30.1452 | 28 | 28 | 30 | 27 | 30 | 26.6 | 123 | 29.2195 | 28 | 28.0297 | 31 | 30.7 | 33.05 | 26 | 26.75 | 26.5 | 29.5 | 26.5
| Std Dev. |  | 4.2564 | 1.7363 | 0.74476 | 0.75277 | 0.8165 | 1.1427 | 1.2151 | 1.0328 | 0.5164 | 12.9254 | 0.75277 | 0.75277 | 0.78342 | 4.1833 | 0.44574 | 0.75277 | 0.24197 | 0.5164 | 0.50067 | 0.48477 | 0.5164 | 1.7475 | 1.0488 | 0.8165 | 1.169
| Rel.Std.Dev. |  | 12.51% | 5.489% | 2.513% | 2.495% | 2.692% | 4.204% | 4.061% | 3.645% | 1.866% | 36.24% | 2.771% | 2.523% | 2.927% | 3.387% | 1.538% | 2.673% | 0.8604% | 1.684% | 1.629% | 1.467% | 2.012% | 6.397% | 3.958% | 2.752% | 4.357%
| PDM3 |  | 17.27% | 9.037% | 2.144% | 3.982% | 4.556% | -6.301% | 3.121% | -2.337% | -4.635% | 22.94% | -6.359% | 2.833% | -7.738% | 325.7% | -0.1284% | -2.912% | -3.06% | 5.705% | 5.935% | 13.92% | -11.53% | -5.842% | -8.657% | 2.258% | -7.508%
| Z-Score (Absolute) |  | 1.5 | 0.69 | 0 | 0.18 | 0.24 | 0.84 | 0.1 | 0.45 | 0.67 | 2.07 | 0.85 | 0.07 | 0.98 | 32.18 | 0.23 | 0.5 | 0.52 | 0.35 | 0.38 | 1.17 | 1.36 | 0.79 | 1.07 | 0.01 | 0.96
| Analytical results for Zr in OREAS 555b (Certified Value 188 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 186.1 | 188.7 | 173 | 183.6 | 201 | 179 | 248.7784 | 186.5 | 195.5 | 189.5 | 186.5 | 207 | 160 | 161 | 191.7687 | 195.7311 | 180 | 186.5492 | 168.6 | 190 | 184.2 | 185.3 | 189 | 168 | 196.7 | 182.5 | 192.4 | 1
| 1 | 2 | 188.9 | 190.3 | 175 | 183.8 | 201 | 185 | 251.8348 | 193 | 185.5 | 196.5 | 184 | 205 | 158 | 156 | 187.3905 | 194.31 | 184 | 185.2662 | 169.6 | 188 | 187.2 | 183.2 | 186 | 195 | 199 | 180.5 | 190.1 | 18
| 1 | 3 | 187.4 | 190.2 | 169 | 187.2 | 199 | 181 | 247.2798 | 196 | 191 | 194 | 190.5 | 205 | 159 | 163 | 189.5216 | 194.2792 | 187 | 179.3096 | 171.2 | 192 | 176.1 | 183.7 | 186 | 171 | 195.6 | 187.9 | 187.2 | 16
| 1 | 4 | 186.8 | 197.4 | 178 | 183.3 | 191 | 185 | 253.1708 | 189 | 183.5 | 198 | 187.5 | 197.5 | 157 | 166 | 193.0901 | 198.647 | 181 | 185.957 | 167.8 | 189 | 181.5 | 179.6 | 173 | 181 | 193.7 | 183.3 | 192.5 | 187.9751
| 1 | 5 | 182.9 | 196.7 | 177 | 187.1 | 195 | 190 | 251.9845 | 189 | 191 | 193.5 | 203 | 205 | 162 | 159 | 193.1188 | 197.725 | 187 | 186.4821 | 166.7 | 193 | 182.1 | 184.7 | 187 | 197 | 199.8 | 184.4 | 193.9 | 146
| 1 | 6 | 183.6 | 191.1 | 178 | 181.1 | 196 | 191 | 252.2574 | 189.5 | 184.5 | 194.5 | 189 | 203 | 163 | 158 | 188.1885 | 196.2092 | 184 | 183.0235 | 160.7 | 195 | 180 | 182.5 | 180 | 184 | 196 | 182.7 | 189.3
| Mean |  | 185.95 | 192.4 | 175 | 184.35 | 197.1667 | 185.1667 | 250.8843 | 190.5 | 188.5 | 194.3333 | 190.0833 | 203.75 | 159.8333 | 160.5 | 190.513 | 196.1503 | 183.8333 | 184.4312 | 167.4333 | 191.1667 | 181.85 | 183.1667 | 183.5 | 182.6667 | 196.8 | 183.55 | 190.9
| Median |  | 186.45 | 190.7 | 176 | 183.7 | 197.5 | 185 | 251.9096 | 189.25 | 188.25 | 194.25 | 188.25 | 205 | 159.5 | 160 | 190.6451 | 195.9702 | 184 | 185.6116 | 168.2 | 191 | 181.8 | 183.45 | 186 | 182.5 | 196.35 | 183 | 191.25
| Std Dev. |  | 2.2967 | 3.6911 | 3.5214 | 2.3755 | 3.92 | 4.7504 | 2.3089 | 3.4059 | 4.7223 | 2.9098 | 6.7039 | 3.3129 | 2.3166 | 3.6194 | 2.4958 | 1.7765 | 2.9269 | 2.8259 | 3.6401 | 2.6394 | 3.7634 | 2.0176 | 5.9582 | 11.9443 | 2.2601 | 2.4833 | 2.4779
| Rel.Std.Dev. |  | 1.235% | 1.918% | 2.012% | 1.289% | 1.988% | 2.565% | 0.9203% | 1.788% | 2.505% | 1.497% | 3.527% | 1.626% | 1.449% | 2.255% | 1.31% | 0.9057% | 1.592% | 1.532% | 2.174% | 1.381% | 2.069% | 1.102% | 3.247% | 6.539% | 1.148% | 1.353% | 1.298%
| PDM3 |  | -1.077% | 2.354% | -6.903% | -1.929% | 4.89% | -1.494% | 33.47% | 1.343% | 0.2792% | 3.382% | 1.122% | 8.392% | -14.97% | -14.62% | 1.35% | 4.349% | -2.203% | -1.885% | -10.93% | 1.698% | -3.258% | -2.558% | -2.381% | -2.824% | 4.695% | -2.354% | 1.556%
| Z-Score (Absolute) |  | 0 | 0.95 | 1.62 | 0.24 | 1.66 | 0.12 | 9.6 | 0.67 | 0.38 | 1.24 | 0.61 | 2.63 | 3.86 | 3.76 | 0.67 | 1.51 | 0.31 | 0.22 | 2.74 | 0.77 | 0.61 | 0.41 | 0.36 | 0.49 | 1.6 | 0.35 | 0.73
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C825: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C862: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1047: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1065: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1084: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1102: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1120: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1138: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 555b (Certified Value 0.076 ppm) | [28 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | 1.3 | 0.098 | < 0.5 | 0.09 | 0.065 | 0.1 | < 0.2 | 0.08 | 0.07 | 0.1 | 0.09 | 0.07 | < 5 | 0.07 | 0.07205 | 0.05 | < 0.1 | < 0.3 | < 0.1 | 0.3 | < 0.1 | 0.09 | 0.08 | < 0.5 | 0.07 |  |  | 1
| 1 | 2 | 1.37 | 0.069 | < 0.5 | 0.09 | 0.071 | 0.09 | < 0.2 | 0.08 | 0.07 | 0.08 | 0.08 | 0.07 | < 5 | 0.07 | 0.055033 | 0.06 | < 0.1 | < 0.3 | < 0.1 | 0.2 | < 0.1 | 0.08 | 0.06 | < 0.5 | 0.09 |  |  | 1
| 1 | 3 | 1.47 | 0.098 | < 0.5 | 0.09 | 0.069 | 0.09 | < 0.2 | 0.08 | 0.07 | 0.08 | 0.08 | 0.07 | < 5 | 0.08 | < 0.05 | 0.05 | < 0.1 | < 0.3 | < 0.1 | 0.4 | < 0.1 | 0.08 | 0.05 | < 0.5 | 0.07 |  |  | 16
| 1 | 4 | 1.22 | 0.078 | < 0.5 | 0.08 | 0.059 | 0.09 | < 0.2 | 0.08 | 0.07 | 0.08 | 0.08 | 0.08 | < 5 | 0.11 | 0.053717 | 0.06 | < 0.1 | < 0.3 | < 0.1 | 0.3 | < 0.1 | 0.08 | 0.06 | < 0.5 | 0.07 |  |  | 0.075622
| 1 | 5 | 1.16 | 0.078 | < 0.5 | 0.09 | 0.071 | 0.09 | < 0.2 | 0.08 | 0.07 | 0.08 | 0.08 | 0.11 | < 5 | 0.08 | 0.051388 | 0.05 | < 0.1 | < 0.3 | < 0.1 | 0.2 | < 0.1 | 0.08 | 0.06 | < 0.5 | 0.07 |  |  | 148
| 1 | 6 | 1.35 | 0.098 | < 0.5 | 0.11 | 0.073 | 0.09 | < 0.2 | 0.08 | 0.07 | 0.08 | 0.06 | 0.08 | < 5 | 0.07 | 0.068633 | 0.06 | < 0.1 | < 0.3 | < 0.1 | 0.4 | < 0.1 | 0.09 | 0.07 | < 0.5 | 0.08
| Mean |  | 1.3117 | 0.0865 |  | 0.091667 | 0.068 | 0.091667 |  | 0.08 | 0.07 | 0.083333 | 0.078333 | 0.08 |  | 0.08 | 0.060164 | 0.055 |  |  |  | 0.3 |  | 0.083333 | 0.063333 |  | 0.075
| Median |  | 1.325 | 0.088 |  | 0.09 | 0.07 | 0.09 |  | 0.08 | 0.07 | 0.08 | 0.08 | 0.075 |  | 0.075 | 0.055033 | 0.055 |  |  |  | 0.3 |  | 0.08 | 0.06 |  | 0.07
| Std Dev. |  | 0.11089 | 0.013019 |  | 0.0098319 | 0.0051769 | 0.0040825 |  | 0 | 0 | 0.008165 | 0.0098319 | 0.015492 |  | 0.015492 | 0.0094592 | 0.0054772 |  |  |  | 0.089443 |  | 0.005164 | 0.010328 |  | 0.0083666
| Rel.Std.Dev. |  | 8.454% | 15.05% |  | 10.73% | 7.613% | 4.454% |  | 0% | 0% | 9.798% | 12.55% | 19.36% |  | 19.36% | 15.72% | 9.959% |  |  |  | 29.81% |  | 6.197% | 16.31% |  | 11.16%
| PDM3 |  | 1635% | 14.38% |  | 21.22% | -10.08% | 21.22% |  | 5.789% | -7.434% | 10.2% | 3.585% | 5.789% |  | 5.789% | -20.44% | -27.27% |  |  |  | 296.7% |  | 10.2% | -16.25% |  | -0.8226%
| Z-Score (Absolute) |  | 57.94 | 0.31 | 8 | 0.55 | 0.56 | 0.55 | 0.94 | 0 | 0.47 | 0.16 | 0.08 | 0 | 113.85 | 0 | 1.21 | 1.18 | 1.41 | 3.29 | 1.41 | NA | 1.41 | 0.16 | 0.78 | 8 | 0.24
Results from laboratory 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 555b (Certified Value 0.959 wt.%) | [28 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.87 | 0.82 | 1.2 | 1.0668 | 1.33 | 0.54 | 0.8 | 1 | 0.85 | 0.83 | 0.88 | 0.8988 | 0.73 | 1.319 | 1.49 | 0.94068 | 0.67 | 0.66 | 1 | 1.207 | 1.51 | 1.506 | 1 | 0.97 |  |  |  | 1
| 1 | 2 | 0.83 | 0.93 | 1.16 | 1.0712 | 1.27 | 0.54 | 0.81 | 1.03 | 0.85 | 0.83 | 0.88 | 0.9147 | 0.693 | 1.1908 | 1.52 | 0.98849 | 0.66 | 0.66 | 1.02 | 1.21 | 1.88 | 1.4948 | 0.94 | 0.94 |  |  |  | #N/A
| 1 | 3 | 0.85 | 1.01 | 1.2 | 1.1037 | 1.28 | 0.53 | 0.84 | 1.03 | 0.82 | 0.85 | 0.88 | 0.9131 | 0.736 | 1.2225 | 1.51 | 0.95832 | 0.67 | 0.65 | 1.02 | 1.189 | 1.78 | 1.4701 | 0.97 | 0.95 |  |  |  | 16
| 1 | 4 | 0.86 | 1.05 | 1.19 | 1.0944 | 1.32 | 0.53 | 0.84 | 1.02 | 0.83 | 0.85 | 0.9 | 0.9125 | 0.651 | 1.1991 | 1.55 | 0.94815 | 0.68 | 0.65 | 1.01 | 1.206 | 1.84 | 1.5468 | 0.99 | 0.93 |  |  |  | 0.9585
| 1 | 5 | 0.88 | 1.13 | 1.22 | 1.1059 | 1.31 | 0.5 | 0.86 | 1.02 | 0.85 | 0.84 | 0.87 | 0.9124 | 0.651 | 1.2842 | 1.51 | 0.96107 | 0.66 | 0.67 | 0.99 | 1.221 | 1.21 | 1.4778 | 1 | 0.96 |  |  |  | 149
| 1 | 6 | 0.87 | 1.07 | 1.21 | 1.0772 | 1.27 | 0.55 | 0.84 | 1.02 | 0.85 | 0.87 | 0.84 | 0.9136 | 0.64 | 1.309 | 1.54 | 0.95417 | 0.66 | 0.66 | 1.02 | 1.246 | 1.25 | 1.519 | 1 | 0.99
| Mean |  | 0.86 | 1.0017 | 1.1967 | 1.0865 | 1.2967 | 0.53167 | 0.83167 | 1.02 | 0.84167 | 0.845 | 0.875 | 0.91085 | 0.6835 | 1.2541 | 1.52 | 0.95848 | 0.66667 | 0.65833 | 1.01 | 1.2132 | 1.5783 | 1.5024 | 0.98333 | 0.95667
| Median |  | 0.865 | 1.03 | 1.2 | 1.0858 | 1.295 | 0.535 | 0.84 | 1.02 | 0.85 | 0.845 | 0.88 | 0.9128 | 0.672 | 1.2533 | 1.515 | 0.95624 | 0.665 | 0.66 | 1.015 | 1.2085 | 1.645 | 1.5004 | 0.995 | 0.955
| Std Dev. |  | 0.017889 | 0.11107 | 0.020656 | 0.01699 | 0.026583 | 0.017224 | 0.022286 | 0.010954 | 0.013292 | 0.015166 | 0.019748 | 0.0059628 | 0.042458 | 0.056857 | 0.021909 | 0.01643 | 0.008165 | 0.0075277 | 0.012649 | 0.019094 | 0.29943 | 0.028164 | 0.024221 | 0.021602
| Rel.Std.Dev. |  | 2.08% | 11.09% | 1.726% | 1.564% | 2.05% | 3.24% | 2.68% | 1.074% | 1.579% | 1.795% | 2.257% | 0.6546% | 6.212% | 4.534% | 1.441% | 1.714% | 1.225% | 1.143% | 1.252% | 1.574% | 18.97% | 1.875% | 2.463% | 2.258%
| PDM3 |  | -10.28% | 4.503% | 24.85% | 13.36% | 35.28% | -44.53% | -13.23% | 6.416% | -12.19% | -11.84% | -8.712% | -4.972% | -28.69% | 30.84% | 58.58% | -0.002265% | -30.45% | -31.32% | 5.373% | 26.57% | 64.67% | 56.75% | 2.591% | -0.1916%
| Z-Score (Absolute) |  | 0.56 | 0.15 | 1.13 | 0.58 | 1.64 | 2.21 | 0.7 | 0.25 | 0.65 | 0.63 | 0.48 | 0.3 | 1.44 | 1.42 | 2.76 | 0.06 | 1.53 | 1.57 | 0.2 | 1.22 | 3.05 | 2.67 | 0.06 | 0.07
| Analytical results for As in OREAS 555b (Certified Value 62 ppm) | [28 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | 1
| 1 | 1 | 42.4 | 50.4 | 61.4 | 63 | 64.6 | 64.5 | 60.8197 | 63.2 | 66.2 | 62.4 | 63.7 | 66.9 | 50.22 | 59.7 | 58.836 | 55.5 | 61 | 62.4986 | 55 | 62.7 | 64 | 60.8 | 51 | 65.4 | 62 | 61 | 63.3 | 1
| 1 | 2 | 45.9 | 57.1 | 61.8 | 63 | 64.2 | 64.9 | 62.3974 | 65.2 | 68.1 | 63.5 | 62.9 | 64.2 | 53.44 | 63.9 | 58.6673 | 54.08 | 61 | 63.4698 | 56 | 64.3 | 60 | 58.6 | 57 | 60.4 | 64 | 60 | 63.6 | 2
| 1 | 3 | 47.2 | 57 | 61.1 | 64 | 62.7 | 64.7 | 59.8921 | 65.6 | 65.4 | 60.9 | 63.8 | 65.3 | 52.48 | 65 | 59.7449 | 56.02 | 61 | 63.5031 | 56.5 | 65.8 | 59 | 56.2 | 55 | 59.7 | 61 | 60 | 63.2 | 16
| 1 | 4 | 46.9 | 58.7 | 60.6 | 65 | 60.7 | 66.6 | 61.9932 | 65.3 | 66.4 | 61.5 | 64.2 | 64.4 | 52.76 | 62 | 59.1258 | 54.57 | 60 | 63.3797 | 54.7 | 63.3 | 57 | 59.5 | 56 | 57.1 | 61 | 62 | 61.8 | 61.6015
| 1 | 5 | 44.7 | 59.5 | 61.9 | 62 | 63.6 | 67.4 | 62.2146 | 65.8 | 66.8 | 67.3 | 64.4 | 65.1 | 52.73 | 62 | 59.8377 | 54.66 | 61 | 64.028 | 57.6 | 62.8 | 62 | 58 | 58 | 56.8 | 62 | 60 | 62.8 | 150
| 1 | 6 | 43 | 60 | 62.6 | 64 | 62.9 | 67.4 | 61.4635 | 65.1 | 65.7 | 63.7 | 63 | 65.7 | 52.8 | 63.2 | 59.2017 | 54.41 | 60 | 63.8516 | 57.9 | 62.2 | 58 | 58.9 | 59 | 59.5 | 63 | 62 | 62.6
| Mean |  | 45.0167 | 57.1167 | 61.5667 | 63.5 | 63.1167 | 65.9167 | 61.4634 | 65.0333 | 66.4333 | 63.2167 | 63.6667 | 65.2667 | 52.405 | 62.6333 | 59.2356 | 54.8733 | 60.6667 | 63.4551 | 56.2833 | 63.5167 | 60 | 58.6667 | 56 | 59.8167 | 62.1667 | 60.8333 | 62.8833
| Median |  | 45.3 | 57.9 | 61.6 | 63.5 | 63.25 | 65.75 | 61.7283 | 65.25 | 66.3 | 62.95 | 63.75 | 65.2 | 52.745 | 62.6 | 59.1637 | 54.615 | 61 | 63.4864 | 56.25 | 63.05 | 59.5 | 58.75 | 56.5 | 59.6 | 62 | 60.5 | 63
| Std Dev. |  | 2.0054 | 3.5108 | 0.69474 | 1.0488 | 1.3906 | 1.3703 | 0.95924 | 0.93524 | 0.95638 | 2.279 | 0.6121 | 0.97707 | 1.117 | 1.8403 | 0.4728 | 0.73337 | 0.5164 | 0.53078 | 1.3136 | 1.3258 | 2.6077 | 1.5384 | 2.8284 | 3.1019 | 1.169 | 0.98319 | 0.64005
| Rel.Std.Dev. |  | 4.455% | 6.147% | 1.128% | 1.652% | 2.203% | 2.079% | 1.561% | 1.438% | 1.44% | 3.605% | 0.9614% | 1.497% | 2.132% | 2.938% | 0.7982% | 1.336% | 0.8512% | 0.8365% | 2.334% | 2.087% | 4.346% | 2.622% | 5.051% | 5.186% | 1.881% | 1.616% | 1.018%
| PDM3 |  | -26.92% | -7.28% | -0.05654% | 3.082% | 2.46% | 7.005% | -0.2241% | 5.571% | 7.844% | 2.622% | 3.352% | 5.95% | -14.93% | 1.675% | -3.841% | -10.92% | -1.518% | 3.009% | -8.633% | 3.109% | -2.6% | -4.764% | -9.093% | -2.897% | 0.9175% | -1.247% | 2.081%
| Z-Score (Absolute) |  | 5.72 | 1.54 | 0 | 0.67 | 0.54 | 1.5 | 0.04 | 1.2 | 1.68 | 0.57 | 0.73 | 1.28 | 3.17 | 0.37 | 0.81 | 2.31 | 0.31 | 0.65 | 1.83 | 0.67 | 0.54 | 1 | 1.92 | 0.61 | 0.21 | 0.25 | 0.46
| Analytical results for Au in OREAS 555b (Certified Value < 0.02 ppm) | [28 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 25 | 26 | 28 | [13 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [13 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.1g | 01g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.0002 | < 0.001 | < 0.0005 | < 0.005 | 0.00811 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.001 | 0.0031 | 0.001 | 0.0026 | [13 empty] | 1
| 1 | 2 | < 0.0002 | < 0.001 | < 0.0005 | < 0.005 | 0.00618 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.001 | 0.0046 | < 0.001 | 0.0012 | [13 empty] | #N/A
| 1 | 3 | < 0.0002 | 0.001 | < 0.0005 | < 0.005 | 0.01755 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.001 | 0.0012 | < 0.001 | 0.0011 | [13 empty] | 16
| 1 | 4 | < 0.0002 | < 0.001 | < 0.0005 | < 0.005 | < 0.003 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.001 | < 0.0005 | < 0.001 | 0.0019 | [13 empty] | < 0.02
| 1 | 5 | < 0.0002 | < 0.001 | < 0.0005 | < 0.005 | < 0.0073 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.001 | 0.0008 | < 0.001 | 0.0008 | [13 empty] | 151
| 1 | 6 | < 0.0002 | < 0.001 | < 0.0005 | < 0.005 | 0.00863 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.001 | 0.0008 | < 0.001 | < 0.0005
| Mean |  |  | 0.001 |  |  | 0.010117 | [6 empty] | 0.0021 | 0.001 | 0.00152
| Median |  |  | 0.001 |  |  | 0.00837 | [6 empty] | 0.0012 | 0.001 | 0.0012
| Std Dev. | [5 empty] | 0.0050659 | [6 empty] | 0.0016912 |  | 0.00072595
| Rel.Std.Dev. | [5 empty] | 50.07% | [6 empty] | 80.53% |  | 47.76%
PDM3
| Z-Score (Absolute) |  | 0.7 | 0.53 | 0.65 | 0.12 | 1.87 | 2.69 | 2.69 | 2.69 | 2.69 | 2.69 | 0.56 | 0.12 | 0.53 | 0.29
| Analytical results for B in OREAS 555b (Certified Value < 10 ppm) | [28 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 18 | 21 | 25 | 26 | 28 | [12 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | [12 empty] | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.25g | 0.1g | 01g | 0.5g | [12 empty] | 1
| 1 | 1 | < 20 | < 10 | 5 | 2.8 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 2.8623 | < 10 | 12 | < 10 | < 10 | [12 empty] | 1
| 1 | 2 | < 20 | < 10 | 5 | 2.9 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3.1544 | < 10 | 13 | < 10 | < 10 | [12 empty] | #N/A
| 1 | 3 | < 20 | < 10 | 5 | 2.8 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 2.9094 | < 10 | 12 | < 10 | < 10 | [12 empty] | 16
| 1 | 4 | < 20 | < 10 | 5 | 2.5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 2.9912 | < 10 | 11 | < 10 | < 10 | [12 empty] | < 10
| 1 | 5 | < 20 | < 10 | 5 | 3.1 | < 10 | < 10 | < 10 | 10 | < 10 | < 10 | 3.0661 | < 10 | < 10 | < 10 | < 10 | [12 empty] | 152
| 1 | 6 | < 20 | < 10 | 5 | 2.7 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 2.9951 | < 10 | < 10 | < 10 | < 10
| Mean |  |  |  | 5 | 2.8 |  |  |  | 10 |  |  | 2.9964 |  | 12
| Median |  |  |  | 5 | 2.8 |  |  |  | 10 |  |  | 2.9931 |  | 12
| Std Dev. |  |  |  | 0 | 0.2 | [6 empty] | 0.10526 |  | 0.8165
| Rel.Std.Dev. |  |  |  | 0% | 7.143% | [6 empty] | 3.513% |  | 6.804%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Results from laboratories 5 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ba in OREAS 555b (Certified Value 57 ppm) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 0
| 1 | 1 | 63.2 | 50.4 | 28.4 | 39 | 17.7 | 48 | 50 | 50 | 50 | 50 | 50 | 56.14 | 9 | 56.6105 | 75 | 53.0825 | 53 | 49 | 64 | 79.9 | 72 | 58 | 53 |  |  |  |  | 1
| 1 | 2 | 59.3 | 54.9 | 28.2 | 48 | 18.4 | 50 | 50 | 60 | 50 | 50 | 50 | 56.4 | 10 | 55.8609 | 75 | 53.8624 | 53 | 48 | 67 | 80.4 | 72 | 58 | 51 |  |  |  |  | 30
| 1 | 3 | 62.7 | 56.6 | 28.8 | 45 | 19.2 | 50 | 50 | 60 | 50 | 50 | 50 | 56.24 | 10 | 58.1431 | 73 | 53.7661 | 53 | 48 | 64 | 78 | 72 | 57 | 50 |  |  |  |  | 16
| 1 | 4 | 60.6 | 59.4 | 30 | 37 | 16.2 | 52 | 50 | 60 | 50 | 50 | 50 | 56.28 | 7 | 54.8984 | 71 | 52.8308 | 52 | 47 | 65 | 79.8 | 73 | 60 | 51 |  |  |  |  | 56.6609
| 1 | 5 | 60.5 | 60.6 | 29.4 | 45 | 18.7 | 51 | 50 | 60 | 40 | 50 | 50 | 56.31 | 7 | 58.6075 | 73 | 52.1897 | 53 | 46 | 68 | 65.8 | 72 | 59 | 53 |  |  |  |  | 153
| 1 | 6 | 64.2 | 60.3 | 30.5 | 40 | 19.8 | 52 | 50 | 60 | 50 | 50 | 50 | 56.35 | 9 | 57.5546 | 71 | 52.655 | 53 | 48 | 65 | 70.2 | 74 | 60 | 53
| Mean |  | 61.75 | 57.0333 | 29.2167 | 42.3333 | 18.3333 | 50.5 | 50 | 58.3333 | 48.3333 | 50 | 50 | 56.2867 | 8.6667 | 56.9458 | 73 | 53.0644 | 52.8333 | 47.6667 | 65.5 | 75.6833 | 72.5 | 58.6667 | 51.8333
| Median |  | 61.65 | 58 | 29.1 | 42.5 | 18.55 | 50.5 | 50 | 60 | 50 | 50 | 50 | 56.295 | 9 | 57.0825 | 73 | 52.9567 | 53 | 48 | 65 | 78.9 | 72 | 58.5 | 52
| Std Dev. |  | 1.8918 | 3.9429 | 0.91305 | 4.274 | 1.2644 | 1.5166 | 0 | 4.0825 | 4.0825 | 0 | 0 | 0.090701 | 1.3663 | 1.4175 | 1.7889 | 0.65066 | 0.40825 | 1.0328 | 1.6432 | 6.1658 | 0.83666 | 1.2111 | 1.3292
| Rel.Std.Dev. |  | 3.064% | 6.913% | 3.125% | 10.1% | 6.897% | 3.003% | 0% | 6.999% | 8.447% | 0% | 0% | 0.1611% | 15.76% | 2.489% | 2.45% | 1.226% | 0.7727% | 2.167% | 2.509% | 8.147% | 1.154% | 2.064% | 2.564%
| PDM3 |  | 8.982% | 0.6573% | -48.44% | -25.29% | -67.64% | -10.87% | -11.76% | 2.952% | -14.7% | -11.76% | -11.76% | -0.6605% | -84.7% | 0.5028% | 28.84% | -6.347% | -6.755% | -15.87% | 15.6% | 33.57% | 27.95% | 3.54% | -8.52%
| Z-Score (Absolute) |  | 1.05 | 0.48 | 2.87 | 1.29 | 4.18 | 0.31 | 0.37 | NA | NA | NA | NA | 0.39 | 5.34 | 0.47 | 2.4 | 0 | 0.03 | 0.65 | 1.5 | 2.72 | 2.34 | 0.67 | 0.15
Results from laboratories 9, 10, 11 and 12 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 555b (Certified Value 2.21 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [5 empty] | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [5 empty] | 2
| 1 | 1 | 2.6 | < 5 | 2.24 | 2 | 2.03 | 2.03 | 2.21 | 2.19 | 2.09 | 2.11 | < 5 | 1.75 | 1.893 | 2.5 | 2.2157 | 2.7 | 1.9 | 2 | 2.01 | 2.04 | 2.71 | 2.45 | [5 empty] | 1
| 1 | 2 | 2.4 | < 5 | 2.31 | 2.2 | 2.06 | 2.07 | 2.28 | 2.2 | 2.06 | 2.09 | < 5 | 1.77 | 1.8812 | 2.5 | 2.2842 | 3 | 1.9 | 3 | 2.22 | 2.1 | 2.69 | 2.4 | [5 empty] | 3
| 1 | 3 | 2.4 | < 5 | 2.18 | 2.1 | 1.84 | 2.09 | 2.27 | 2.18 | 2.09 | 2.11 | < 5 | 1.84 | 2.0857 | 2.5 | 2.2699 | 3 | 2 | 3 | 2.1 | 2.11 | 2.64 | 2.39 | [5 empty] | 16
| 1 | 4 | 2.5 | < 5 | 2.27 | 2 | 1.83 | 2.12 | 2.29 | 2.2 | 2.1 | 2.11 | < 5 | 1.8 | 2.0419 | 2.5 | 2.2431 | 2.5 | 1.9 | 2 | 2.22 | 2.02 | 2.76 | 2.39 | [5 empty] | 2.2126
| 1 | 5 | 2.5 | < 5 | 2.44 | 2.1 | 1.99 | 2.14 | 2.27 | 2.21 | 2.08 | 2.09 | < 5 | 1.79 | 2.066 | 2.5 | 2.2139 | 2.6 | 2 | 3 | 2.26 | 1.66 | 2.74 | 2.42 | [5 empty] | 154
| 1 | 6 | 2.6 | < 5 | 2.33 | 2.1 | 2.07 | 2.09 | 2.25 | 2.21 | 2.11 | 2.04 | < 5 | 1.81 | 1.9823 | 2.4 | 2.2317 | 2.8 | 2 | 3 | 2.29 | 1.65 | 2.7 | 2.43
| Mean |  | 2.5 |  | 2.295 | 2.0833 | 1.97 | 2.09 | 2.2617 | 2.1983 | 2.0883 | 2.0917 |  | 1.7933 | 1.9917 | 2.4833 | 2.2431 | 2.7667 | 1.95 | 2.6667 | 2.1833 | 1.93 | 2.7067 | 2.4133
| Median |  | 2.5 |  | 2.29 | 2.1 | 2.01 | 2.09 | 2.27 | 2.2 | 2.09 | 2.1 |  | 1.795 | 2.0121 | 2.5 | 2.2374 | 2.75 | 1.95 | 3 | 2.22 | 2.03 | 2.705 | 2.41
| Std Dev. |  | 0.089443 |  | 0.088713 | 0.075277 | 0.10826 | 0.038471 | 0.028577 | 0.01169 | 0.017224 | 0.027142 |  | 0.031411 | 0.088204 | 0.040825 | 0.028778 | 0.20656 | 0.054772 | 0.5164 | 0.10671 | 0.21578 | 0.041793 | 0.024221
| Rel.Std.Dev. |  | 3.578% |  | 3.865% | 3.613% | 5.495% | 1.841% | 1.264% | 0.5318% | 0.8248% | 1.298% |  | 1.752% | 4.429% | 1.644% | 1.283% | 7.466% | 2.809% | 19.36% | 4.887% | 11.18% | 1.544% | 1.004%
| PDM3 |  | 12.99% |  | 3.723% | -5.843% | -10.97% | -5.542% | 2.217% | -0.6456% | -5.617% | -5.466% |  | -18.95% | -9.986% | 12.24% | 1.377% | 25.04% | -11.87% | 20.52% | -1.324% | -12.77% | 22.33% | 9.071%
| Z-Score (Absolute) |  | 0.95 | 0.95 | 0.3 | 0.36 | 0.72 | 0.34 | 0.2 | 0 | 0.34 | 0.33 | 0.95 | 1.27 | 0.65 | 0.89 | 0.14 | 1.78 | 0.78 | NA | 0.05 | 0.84 | 1.59 | 0.67
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 555b (Certified Value 1.39 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | 3
| 1 | 1 | < 5 | 1.3 | < 2 | 1.54 | 1.35 | 1.42 | 1.3896 | 1.35 | 1.36 | 1.4 | 1.44 | 1.35 | < 5 | 1.2 | 1.3457 | 0.89452 | 1.41 | 1.3072 | 1.37 | 1.46 | < 5 | 1.1 | 1.57 | 1.33 | 1.47 | 1.4 | 1.45 | 1
| 1 | 2 | < 5 | 1.24 | < 2 | 1.52 | 1.31 | 1.48 | 1.3835 | 1.36 | 1.36 | 1.38 | 1.47 | 1.36 | < 5 | 1.29 | 1.3453 | 0.87779 | 1.43 | 1.3726 | 1.24 | 1.45 | < 5 | 1.1 | 1.56 | 1.28 | 1.49 | 1.5 | 1.42 | 4
| 1 | 3 | < 5 | 1.35 | < 2 | 1.47 | 1.28 | 1.45 | 1.3746 | 1.38 | 1.4 | 1.38 | 1.4 | 1.34 | < 5 | 1.33 | 1.3661 | 0.9561 | 1.43 | 1.2768 | 1.3 | 1.48 | < 5 | 1.1 | 1.48 | 1.28 | 1.43 | 1.4 | 1.41 | 16
| 1 | 4 | < 5 | 1.34 | < 2 | 1.49 | 1.26 | 1.47 | 1.4057 | 1.35 | 1.39 | 1.36 | 1.44 | 1.34 | < 5 | 1.27 | 1.3683 | 0.94831 | 1.42 | 1.2645 | 1.3 | 1.46 | 5 | 1.1 | 1.51 | 1.29 | 1.44 | 1.4 | 1.4 | 1.3874
| 1 | 5 | < 5 | 1.43 | < 2 | 1.47 | 1.32 | 1.45 | 1.4676 | 1.38 | 1.41 | 1.42 | 1.47 | 1.32 | < 5 | 1.27 | 1.3616 | 0.89832 | 1.43 | 1.4158 | 1.33 | 1.46 | 7 | 1.1 | 1.59 | 1.29 | 1.48 | 1.4 | 1.4 | 155
| 1 | 6 | < 5 | 1.34 | < 2 | 1.49 | 1.32 | 1.51 | 1.4393 | 1.36 | 1.4 | 1.38 | 1.45 | 1.31 | < 5 | 1.24 | 1.3558 | 0.87451 | 1.41 | 1.3262 | 1.35 | 1.43 | 5 | 1.1 | 1.56 | 1.3 | 1.49 | 1.5 | 1.37
| Mean |  |  | 1.3333 |  | 1.4967 | 1.3067 | 1.4633 | 1.41 | 1.3633 | 1.3867 | 1.3867 | 1.445 | 1.3367 |  | 1.2667 | 1.3571 | 0.90826 | 1.4217 | 1.3272 | 1.315 | 1.4567 | 5.6667 | 1.1 | 1.545 | 1.295 | 1.4667 | 1.4333 | 1.4083
| Median |  |  | 1.34 |  | 1.49 | 1.315 | 1.46 | 1.3976 | 1.36 | 1.395 | 1.38 | 1.445 | 1.34 |  | 1.27 | 1.3587 | 0.89642 | 1.425 | 1.3167 | 1.315 | 1.46 | 5 | 1.1 | 1.56 | 1.29 | 1.475 | 1.4 | 1.405
| Std Dev. |  |  | 0.062503 |  | 0.028048 | 0.032042 | 0.030768 | 0.036261 | 0.013663 | 0.021602 | 0.020656 | 0.025884 | 0.018619 |  | 0.044121 | 0.0099783 | 0.035349 | 0.0098319 | 0.057937 | 0.045935 | 0.01633 | 1.1547 | 2.4324e-16 | 0.041352 | 0.018708 | 0.02582 | 0.05164 | 0.026394
| Rel.Std.Dev. |  |  | 4.688% |  | 1.874% | 2.452% | 2.103% | 2.572% | 1.002% | 1.558% | 1.49% | 1.791% | 1.393% |  | 3.483% | 0.7353% | 3.892% | 0.6916% | 4.365% | 3.493% | 1.121% | 20.38% | 2.211e-14% | 2.677% | 1.445% | 1.76% | 3.603% | 1.874%
| PDM3 |  |  | -3.895% |  | 7.878% | -5.817% | 5.476% | 1.633% | -1.732% | -0.05051% | -0.05051% | 4.154% | -3.654% |  | -8.7% | -2.18% | -34.53% | 2.472% | -4.337% | -5.216% | 4.995% | 308.4% | -20.71% | 11.36% | -6.658% | 5.716% | 3.313% | 1.511%
| Z-Score (Absolute) |  | 10.72 | 0.51 | 3.72 | 1.06 | 0.77 | 0.74 | 0.23 | 0.22 | 0 | 0 | 0.56 | 0.48 | 10.72 | 1.16 | 0.28 | 4.61 | 0.34 | 0.57 | 0.69 | 0.67 | NA | 2.76 | 1.53 | 0.88 | NA | 0.45 | 0.21
Results from laboratories 21 and 26 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 555b (Certified Value 0.137 wt.%) | [28 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | 0.16 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.1356 | 0.143 | 0.13799 | 0.13 | 0.12904 | 0.1478 | 0.14 | 0.13 | 0.14 | 0.145 | 0.128 | 0.14 | 0.13 | 0.14 |  |  | 1
| 1 | 2 | 0.15 | 0.14 | 0.13 | 0.13 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.1344 | 0.143 | 0.12447 | 0.13 | 0.13466 | 0.1486 | 0.13 | 0.13 | 0.14 | 0.151 | 0.123 | 0.14 | 0.13 | 0.14 |  |  | #N/A
| 1 | 3 | 0.15 | 0.14 | 0.13 | 0.14 | 0.12 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.1362 | 0.15 | 0.1339 | 0.13 | 0.13087 | 0.1481 | 0.13 | 0.13 | 0.14 | 0.142 | 0.119 | 0.14 | 0.13 | 0.14 |  |  | 16
| 1 | 4 | 0.14 | 0.14 | 0.13 | 0.14 | 0.12 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.1374 | 0.136 | 0.13106 | 0.13 | 0.13214 | 0.1486 | 0.13 | 0.13 | 0.14 | 0.148 | 0.129 | 0.14 | 0.13 | 0.14 |  |  | 0.13664
| 1 | 5 | 0.15 | 0.15 | 0.14 | 0.13 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.1368 | 0.136 | 0.12232 | 0.13 | 0.13644 | 0.1481 | 0.13 | 0.13 | 0.14 | 0.15 | 0.121 | 0.14 | 0.13 | 0.14 |  |  | 156
| 1 | 6 | 0.16 | 0.15 | 0.14 | 0.13 | 0.12 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.13 | 0.1366 | 0.136 | 0.13789 | 0.13 | 0.13174 | 0.148 | 0.13 | 0.13 | 0.14 | 0.157 | 0.126 | 0.14 | 0.14 | 0.14
| Mean |  | 0.15167 | 0.14167 | 0.135 | 0.13333 | 0.125 | 0.13167 | 0.13833 | 0.14 | 0.14 | 0.14 | 0.13833 | 0.13617 | 0.14067 | 0.13127 | 0.13 | 0.13248 | 0.1482 | 0.13167 | 0.13 | 0.14 | 0.14883 | 0.12433 | 0.14 | 0.13167 | 0.14
| Median |  | 0.15 | 0.14 | 0.135 | 0.13 | 0.125 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.1364 | 0.1395 | 0.13248 | 0.13 | 0.13194 | 0.1481 | 0.13 | 0.13 | 0.14 | 0.149 | 0.1245 | 0.14 | 0.13 | 0.14
| Std Dev. |  | 0.0075277 | 0.0075277 | 0.0054772 | 0.005164 | 0.0054772 | 0.0040825 | 0.0040825 | 0 | 0 | 0 | 0.0040825 | 0.0010539 | 0.0057155 | 0.0066661 | 0 | 0.0026642 | 0.00032863 | 0.0040825 | 0 | 0 | 0.0051929 | 0.0039833 | 0 | 0.0040825 | 0
| Rel.Std.Dev. |  | 4.963% | 5.314% | 4.057% | 3.873% | 4.382% | 3.101% | 2.951% | 0% | 0% | 0% | 2.951% | 0.774% | 4.063% | 5.078% | 0% | 2.011% | 0.2218% | 3.101% | 0% | 0% | 3.489% | 3.204% | 0% | 3.101% | 0%
| PDM3 |  | 11% | 3.679% | -1.2% | -2.42% | -8.519% | -3.64% | 1.239% | 2.459% | 2.459% | 2.459% | 1.239% | -0.3466% | 2.947% | -3.929% | -4.86% | -3.045% | 8.46% | -3.64% | -4.86% | 2.459% | 8.923% | -9.007% | 2.459% | -3.64% | 2.459%
| Z-Score (Absolute) |  | 1.8 | 0.45 | 0.45 | 0.67 | 1.8 | 0.9 | 0 | 0.22 | 0.22 | 0.22 | 0 | 0.29 | 0.3 | 0.95 | 1.12 | 0.79 | 1.33 | 0.9 | 1.12 | 0.22 | 1.42 | 1.89 | 0.22 | 0.9 | 0.22
| Analytical results for Cd in OREAS 555b (Certified Value 0.46 ppm) | [28 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | < 1 | 0.47 | < 0.5 | 0.46 | 0.53 | 0.45 | 0.50086 | 0.46 | 0.44 | 0.47 | 0.46 | 0.47 | < 5 | 0.49 | 0.35432 | 0.39 | < 0.5 | 0.47 | 1 | 0.4 | 0.45 | 0.5 | < 0.5 | 0.47 |  |  |  | 1
| 1 | 2 | < 1 | 0.43 | < 0.5 | 0.46 | 0.46 | 0.47 | 0.42445 | 0.49 | 0.45 | 0.46 | 0.48 | 0.45 | < 5 | 0.51 | 0.33971 | 0.41 | < 0.5 | 0.47 | 1 | 0.4 | 0.46 | 0.47 | < 0.5 | 0.47 |  |  |  | 5
| 1 | 3 | < 1 | 0.49 | < 0.5 | 0.42 | 0.44 | 0.47 | 0.48931 | 0.45 | 0.46 | 0.46 | 0.42 | 0.45 | < 5 | 0.53 | 0.34269 | 0.44 | < 0.5 | 0.47 | 1 | 0.4 | 0.46 | 0.49 | < 0.5 | 0.47 |  |  |  | 16
| 1 | 4 | < 1 | 0.4 | < 0.5 | 0.49 | 0.5 | 0.51 | 0.41546 | 0.45 | 0.46 | 0.45 | 0.45 | 0.49 | < 5 | 0.51 | 0.34327 | 0.41 | < 0.5 | 0.45 | 1 | 0.4 | 0.43 | 0.5 | < 0.5 | 0.47 |  |  |  | 0.46427
| 1 | 5 | < 1 | 0.44 | < 0.5 | 0.41 | 0.48 | 0.51 | 0.45178 | 0.45 | 0.46 | 0.48 | 0.48 | 0.45 | < 5 | 0.52 | 0.3585 | 0.4 | < 0.5 | 0.46 | 1 | 0.4 | 0.46 | 0.46 | < 0.5 | 0.49 |  |  |  | 157
| 1 | 6 | < 1 | 0.48 | < 0.5 | 0.41 | 0.53 | 0.49 | 0.48718 | 0.47 | 0.46 | 0.46 | 0.46 | 0.46 | < 5 | 0.52 | 0.36207 | 0.39 | < 0.5 | 0.46 | 1 | 0.4 | 0.44 | 0.5 | < 0.5 | 0.46
| Mean |  |  | 0.45167 |  | 0.44167 | 0.49 | 0.48333 | 0.46151 | 0.46167 | 0.455 | 0.46333 | 0.45833 | 0.46167 |  | 0.51333 | 0.35009 | 0.40667 |  | 0.46333 | 1 | 0.4 | 0.45 | 0.48667 |  | 0.47167
| Median |  |  | 0.455 |  | 0.44 | 0.49 | 0.48 | 0.46948 | 0.455 | 0.46 | 0.46 | 0.46 | 0.455 |  | 0.515 | 0.3488 | 0.405 |  | 0.465 | 1 | 0.4 | 0.455 | 0.495 |  | 0.47
| Std Dev. |  |  | 0.034303 |  | 0.033116 | 0.036878 | 0.024221 | 0.036246 | 0.016021 | 0.0083666 | 0.010328 | 0.022286 | 0.016021 |  | 0.013663 | 0.0093945 | 0.018619 |  | 0.008165 | 0 | 6.0809e-17 | 0.012649 | 0.017512 |  | 0.0098319
| Rel.Std.Dev. |  |  | 7.595% |  | 7.498% | 7.526% | 5.011% | 7.854% | 3.47% | 1.839% | 2.229% | 4.862% | 3.47% |  | 2.662% | 2.683% | 4.578% |  | 1.762% | 0% | 1.52e-14% | 2.811% | 3.598% |  | 2.085%
| PDM3 |  |  | -2.716% |  | -4.869% | 5.541% | 4.105% | -0.5961% | -0.5617% | -1.998% | -0.2027% | -1.28% | -0.5617% |  | 10.57% | -24.59% | -12.41% |  | -0.2027% | 115.4% | -13.84% | -3.075% | 4.823% |  | 1.592%
| Z-Score (Absolute) |  | 1.24 | 0.32 | 6.85 | 0.64 | 0.92 | 0.7 | 0 | 0 | 0.21 | 0.06 | 0.11 | 0 | 65.98 | 1.67 | 3.61 | 1.78 | 6.85 | 0.06 | NA | NA | 0.38 | 0.81 | 6.85 | 0.33
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution. Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 555b (Certified Value 43.7 ppm) | [28 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | [7 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [7 empty] | 1
| 1 | 1 | 49.391 | 59.8 | 35.1 | 40.5 | 49.7 | 41.7 | 40.9 | 44.4 | 42.95 | 38.4 | 49.3088 | 24.3173 | 54.7 | 43.985 | 33.2 | 41 | 38.56 | 77 | 55 | 43.24 | [7 empty] | 1
| 1 | 2 | 49.267 | 57.1 | 35.8 | 41.8 | 51.3 | 42.8 | 41.1 | 44.5 | 42.79 | 37.2 | 46.3872 | 24.1159 | 52.9 | 44.5408 | 33.2 | 41 | 42.6 | 84.7 | 56 | 42.42 | [7 empty] | 6
| 1 | 3 | 49.669 | 56 | 35.6 | 41.7 | 50.8 | 41.2 | 40.8 | 43.6 | 42.88 | 38.1 | 51.4388 | 24.1975 | 56.1 | 45.1596 | 33.4 | 42 | 42.56 | 85.2 | 56 | 42.77 | [7 empty] | 16
| 1 | 4 | 49.29 | 53.1 | 37 | 42.2 | 50.9 | 43.1 | 41.6 | 44.3 | 42.66 | 35.5 | 47.9231 | 24.9762 | 53.4 | 41.0164 | 34.3 | 41 | 43.65 | 81.9 | 58 | 41.1 | [7 empty] | 43.6838
| 1 | 5 | 49.669 | 56.5 | 36.1 | 42.9 | 51 | 42.9 | 41.2 | 43.5 | 42.78 | 35.3 | 54.53 | 23.5586 | 53.8 | 42.83 | 34.8 | 41 | 45.57 | 61.2 | 55 | 42.95 | [7 empty] | 158
| 1 | 6 | 50.114 | 53.7 | 37 | 42.7 | 49.8 | 41.8 | 42 | 42.9 | 42.83 | 34.5 | 53.6203 | 23.5933 | 52.3 | 44.0023 | 32.9 | 41 | 44.51 | 62.8 | 58 | 43.05
| Mean |  | 49.5667 | 56.0333 | 36.1 | 41.9667 | 50.5833 | 42.25 | 41.2667 | 43.8667 | 42.815 | 36.5 | 50.5347 | 24.1265 | 53.8667 | 43.589 | 33.6333 | 41.1667 | 42.9083 | 75.4667 | 56.3333 | 42.5883
| Median |  | 49.53 | 56.25 | 35.95 | 42 | 50.85 | 42.3 | 41.15 | 43.95 | 42.81 | 36.35 | 50.3738 | 24.1567 | 53.6 | 43.9937 | 33.3 | 41 | 43.125 | 79.45 | 56 | 42.86
| Std Dev. |  | 0.32178 | 2.4328 | 0.76942 | 0.86178 | 0.66758 | 0.78166 | 0.45461 | 0.63456 | 0.098539 | 1.6186 | 3.2195 | 0.52351 | 1.3633 | 1.4761 | 0.74476 | 0.40825 | 2.4217 | 10.8417 | 1.3663 | 0.7804
| Rel.Std.Dev. |  | 0.6492% | 4.342% | 2.131% | 2.053% | 1.32% | 1.85% | 1.102% | 1.447% | 0.2302% | 4.435% | 6.371% | 2.17% | 2.531% | 3.386% | 2.214% | 0.9917% | 5.644% | 14.37% | 2.425% | 1.832%
| PDM3 |  | 13.47% | 28.27% | -17.36% | -3.931% | 15.79% | -3.282% | -5.533% | 0.4185% | -1.989% | -16.45% | 15.68% | -44.77% | 23.31% | -0.217% | -23.01% | -5.762% | -1.775% | 72.76% | 28.96% | -2.508%
| Z-Score (Absolute) |  | 0.69 | 1.36 | 0.7 | 0.09 | 0.8 | 0.06 | 0.16 | 0.1 | 0 | 0.66 | 0.79 | 1.93 | 1.14 | 0.08 | 0.95 | 0.17 | 0 | 3.37 | 1.39 | 0.03
| Analytical results for Co in OREAS 555b (Certified Value 1.08 wt.%) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | AR*AAS | AR*AAS | AR*AAS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/AAS | AR*MS | AR*MS | AR*OES/AAS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*AAS | AR*AAS | AR*OES/MS | AR*MS | AR*OES/MS | Various [a] | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | see footer | 01g | 0.4g |  | 3
| 1 | 1 | 1.07 | 1.13 | 1.1355 | 1.0456 | 1.09 | 1.05 | 1.0499 | 0.916 | 1.085 | 0.944 | 0.931 |  | 1.0626 | 1.02 | 1.1099 | > 0.5 | 1.0938 | 1.0283 | 1.11 | 1.12 | > 1 | > 1 | 1.07 | 1.1377 | 1.0703 | 1.076 |  | 1
| 1 | 2 | 1.08 | 1.11 | 1.1137 | 1.0609 | 1.08 | 1.06 | 1.0107 | 0.937 | 1.06 | 0.941 | 0.935 |  | 1.0622 | 1.06 | 1.0881 | > 0.5 | 1.1328 | 1.0379 | 1.08 | 1.11 | > 1 | > 1 | 1.03 | 1.0633 | 1.045 | 1.095 |  | 36
| 1 | 3 | 1.07 | 1.14 | 1.1325 | 1.1017 | 1.1 | 1.07 | 0.98075 | 0.956 | 1.055 | 0.954 | 0.944 |  | 1.0647 | 1.14 | 1.1009 | > 0.5 | 1.0489 | 1.0311 | 1.14 | 1.1 | > 1 | > 1 | 1.05 | 1.0933 | 1.0464 | 1.132 |  | 16
| 1 | 4 | 1.07 | 1.13 | 1.1229 | 1.0855 | 1.08 | 1.08 | 0.93965 | 0.974 | 1.05 | 0.941 | 0.941 | 1.055 | 1.0651 | 0.967 | 1.1293 | > 0.5 | 1.0959 | 1.0445 | 1.1 | 1.11 | > 1 | > 1 | 1.01 | 1.0873 | 1.0586 | 1.067 |  | 1.0756
| 1 | 5 | 1.08 | 1.14 | 1.1493 | 1.0781 | 1.1 | 1.04 | 0.9814 | 0.979 | 1.085 | 0.973 | 0.957 |  | 1.0655 | 0.992 | 1.0808 | > 0.5 | 1.0821 | 1.0553 | 1.09 | 1.12 | > 1 | > 1 | 1.05 | 1.0927 | 1.0591 | 1.066 |  | 159
| 1 | 6 | 1.07 | 1.11 | 1.1373 | 1.0532 | 1.08 | 1.08 | 0.98876 | 0.96 | 1.06 | 0.961 | 0.949 |  | 1.0659 | 1.04 | 1.1371 | > 0.5 | 1.1492 | 1.034 | 1.1 | 1.1 | > 1 | > 1 | 1.03 | > 0.5 | 1.044 | 1.077
| Mean |  | 1.0733 | 1.1267 | 1.1319 | 1.0708 | 1.0883 | 1.0633 | 0.99186 | 0.95367 | 1.0658 | 0.95233 | 0.94283 | 1.055 | 1.0643 | 1.0365 | 1.1077 |  | 1.1005 | 1.0385 | 1.1033 | 1.11 |  |  | 1.04 | 1.0949 | 1.0539 | 1.0855
| Median |  | 1.07 | 1.13 | 1.134 | 1.0695 | 1.085 | 1.065 | 0.98508 | 0.958 | 1.06 | 0.949 | 0.9425 | 1.055 | 1.0649 | 1.03 | 1.1054 |  | 1.0949 | 1.036 | 1.1 | 1.11 |  |  | 1.04 | 1.0927 | 1.0525 | 1.0765
| Std Dev. |  | 0.005164 | 0.013663 | 0.012308 | 0.021293 | 0.0098319 | 0.01633 | 0.036574 | 0.023653 | 0.015303 | 0.012894 | 0.0094322 |  | 0.0015552 | 0.060596 | 0.022301 |  | 0.03601 | 0.0099765 | 0.020656 | 0.0089443 |  |  | 0.020976 | 0.026905 | 0.010514 | 0.02505
| Rel.Std.Dev. |  | 0.4811% | 1.213% | 1.087% | 1.988% | 0.9034% | 1.536% | 3.687% | 2.48% | 1.436% | 1.354% | 1% |  | 0.1461% | 5.846% | 2.013% |  | 3.272% | 0.9606% | 1.872% | 0.8058% |  |  | 2.017% | 2.457% | 0.9977% | 2.308%
| PDM3 |  | -0.2134% | 4.745% | 5.228% | -0.4458% | 1.181% | -1.143% | -7.788% | -11.34% | -0.9107% | -11.46% | -12.35% | -1.918% | -1.05% | -3.638% | 2.979% |  | 2.308% | -3.45% | 2.576% | 3.195% |  |  | -3.312% | 1.788% | -2.02% | 0.9177%
| Z-Score (Absolute) |  | 0.17 | 1.41 | 1.53 | 0.12 | 0.52 | 0.06 | 1.72 | 2.61 | 0 | 2.64 | 2.86 | 0.25 | 0.03 | 0.68 | 0.97 | NA | 0.8 | 0.63 | 0.87 | 1.03 | NA | NA | 0.6 | 0.67 | 0.28 | 0.46
[a]: Lab 26: Rnd 01 - AR*MS (01g), Rnd 01 - AR*OES (01g)
| Analytical results for Cr in OREAS 555b (Certified Value 35.2 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 1
| 1 | 1 | 31.6 | 37.2 | 38 | 39 | 28.5 | 33 | 35 | 34 | 32 | 34 | 32.77 | 35 | 38.652 | 46 | 32.1491 | 45 | 33 | 31 | 33 | 43 | 43 | 35.8 | 36 |  |  |  |  | 1
| 1 | 2 | 32.5 | 36.5 | 37 | 37 | 28.8 | 34 | 36 | 34 | 32 | 33 | 33.06 | 35 | 42.4363 | 45 | 32.9049 | 46 | 33 | 32 | 38 | 46 | 43 | 35.5 | 35 |  |  |  |  | 37
| 1 | 3 | 33.4 | 36.6 | 38 | 36 | 28.5 | 34 | 36 | 33 | 32 | 34 | 32.81 | 36 | 40.7812 | 46 | 31.987 | 47 | 33 | 31 | 35 | 45 | 43 | 35.2 | 35 |  |  |  |  | 16
| 1 | 4 | 33.7 | 36.7 | 38 | 37 | 28.2 | 34 | 36 | 34 | 32 | 34 | 32.83 | 33 | 38.8388 | 44 | 32.1195 | 46 | 32 | 32 | 36 | 46 | 44 | 36.1 | 35 |  |  |  |  | 35.2455
| 1 | 5 | 36.1 | 37.3 | 36 | 38 | 27.3 | 35 | 36 | 34 | 32 | 33 | 32.94 | 31 | 41.4012 | 45 | 32.8598 | 46 | 33 | 31 | 37 | 37 | 43 | 36.1 | 36 |  |  |  |  | 160
| 1 | 6 | 35.3 | 37.5 | 37 | 37 | 29.1 | 34 | 36 | 34 | 32 | 32 | 32.96 | 33 | 40.3085 | 44 | 32.0256 | 46 | 33 | 31 | 38 | 41 | 44 | 35.6 | 36
| Mean |  | 33.7667 | 36.9667 | 37.3333 | 37.3333 | 28.4 | 34 | 35.8333 | 33.8333 | 32 | 33.3333 | 32.895 | 33.8333 | 40.403 | 45 | 32.341 | 46 | 32.8333 | 31.3333 | 36.1667 | 43 | 43.3333 | 35.7167 | 35.5
| Median |  | 33.55 | 36.95 | 37.5 | 37 | 28.5 | 34 | 36 | 34 | 32 | 33.5 | 32.885 | 34 | 40.5448 | 45 | 32.1343 | 46 | 33 | 31 | 36.5 | 44 | 43 | 35.7 | 35.5
| Std Dev. |  | 1.6872 | 0.41793 | 0.8165 | 1.0328 | 0.61968 | 0.63246 | 0.40825 | 0.40825 | 0 | 0.8165 | 0.11005 | 1.8348 | 1.4694 | 0.89443 | 0.42375 | 0.63246 | 0.40825 | 0.5164 | 1.9408 | 3.5214 | 0.5164 | 0.35449 | 0.54772
| Rel.Std.Dev. |  | 4.997% | 1.131% | 2.187% | 2.766% | 2.182% | 1.86% | 1.139% | 1.207% | 0% | 2.449% | 0.3345% | 5.423% | 3.637% | 1.988% | 1.31% | 1.375% | 1.243% | 1.648% | 5.366% | 8.189% | 1.192% | 0.9925% | 1.543%
| PDM3 |  | -4.196% | 4.883% | 5.924% | 5.924% | -19.42% | -3.534% | 1.668% | -4.007% | -9.208% | -5.425% | -6.669% | -4.007% | 14.63% | 27.68% | -8.241% | 30.51% | -6.844% | -11.1% | 2.614% | 22% | 22.95% | 1.337% | 0.7221%
| Z-Score (Absolute) |  | 0.54 | 0.46 | 0.57 | 0.57 | 2.21 | 0.47 | 0.1 | 0.52 | 1.09 | 0.67 | 0.81 | 0.52 | 1.53 | 2.96 | 0.98 | 3.27 | 0.83 | 1.3 | 0.21 | 2.33 | 2.44 | 0.07 | 0
| Analytical results for Cs in OREAS 555b (Certified Value 0.32 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 25 | 26 | 28 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [8 empty] | 2
| 1 | 1 | 0.42 | 0.38 | 0.28 | 0.19 | 0.23 | 0.23 | 0.25 | 0.2 | 0.16 | 0.44331 | 0.34661 | 0.5 | 0.30997 | 0.34 | 0.2 | 0.3 | 0.7 | 0.53 | 0.35 | [8 empty] | 1
| 1 | 2 | 0.42 | 0.35 | 0.29 | 0.2 | 0.25 | 0.23 | 0.26 | 0.22 | 0.17 | 0.41642 | 0.30602 | 0.54 | 0.31124 | 0.35 | 0.2 | 0.34 | 0.8 | 0.53 | 0.33 | [8 empty] | 9
| 1 | 3 | 0.4 | 0.35 | 0.29 | 0.2 | 0.24 | 0.22 | 0.25 | 0.21 | 0.18 | 0.44155 | 0.35409 | 0.54 | 0.30924 | 0.36 | 0.2 | 0.36 | 0.8 | 0.52 | 0.32 | [8 empty] | 16
| 1 | 4 | 0.37 | 0.38 | 0.29 | 0.2 | 0.24 | 0.23 | 0.26 | 0.22 | 0.17 | 0.44285 | 0.3484 | 0.5 | 0.30194 | 0.36 | 0.2 | 0.35 | 0.7 | 0.52 | 0.3 | [8 empty] | 0.32386
| 1 | 5 | 0.4 | 0.37 | 0.3 | 0.22 | 0.24 | 0.22 | 0.24 | 0.2 | 0.17 | 0.4441 | 0.33281 | 0.52 | 0.30187 | 0.38 | 0.2 | 0.39 | 0.6 | 0.54 | 0.31 | [8 empty] | 161
| 1 | 6 | 0.4 | 0.36 | 0.31 | 0.2 | 0.24 | 0.21 | 0.24 | 0.2 | 0.16 | 0.41744 | 0.31644 | 0.5 | 0.30931 | 0.37 | 0.2 | 0.36 | 0.6 | 0.54 | 0.32
| Mean |  | 0.40167 | 0.365 | 0.29333 | 0.20167 | 0.24 | 0.22333 | 0.25 | 0.20833 | 0.16833 | 0.43428 | 0.33406 | 0.51667 | 0.30726 | 0.36 | 0.2 | 0.35 | 0.7 | 0.53 | 0.32167
| Median |  | 0.4 | 0.365 | 0.29 | 0.2 | 0.24 | 0.225 | 0.25 | 0.205 | 0.17 | 0.4422 | 0.33971 | 0.51 | 0.30927 | 0.36 | 0.2 | 0.355 | 0.7 | 0.53 | 0.32
| Std Dev. |  | 0.018348 | 0.013784 | 0.010328 | 0.0098319 | 0.0063246 | 0.008165 | 0.0089443 | 0.0098319 | 0.0075277 | 0.013469 | 0.0193 | 0.019664 | 0.0042112 | 0.014142 | 3.0405e-17 | 0.029665 | 0.089443 | 0.0089443 | 0.017224
| Rel.Std.Dev. |  | 4.568% | 3.776% | 3.521% | 4.875% | 2.635% | 3.656% | 3.578% | 4.719% | 4.472% | 3.102% | 5.777% | 3.806% | 1.371% | 3.928% | 1.52e-14% | 8.476% | 12.78% | 1.688% | 5.355%
| PDM3 |  | 24.03% | 12.7% | -9.426% | -37.73% | -25.89% | -31.04% | -22.81% | -35.67% | -48.02% | 34.09% | 3.15% | 59.53% | -5.125% | 11.16% | -38.24% | 8.072% | 116.1% | 63.65% | -0.6769%
| Z-Score (Absolute) |  | 0.67 | 0.37 | 0.24 | 1.01 | 0.69 | 0.83 | 0.6 | 0.96 | 1.29 | 0.95 | 0.1 | 1.64 | 0.12 | 0.32 | NA | 0.24 | NA | 1.76 | 0
Results from laboratories 22 and 25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 555b (Certified Value 4.57 wt.%) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | AR*AAS | AR*AAS | AR*AAS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*AAS | AR*AAS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.4g |  | 3
| 1 | 1 | 4.47 | 4.561 | 4.55 | 4.6154 | 4.62 | 4.54 | 4.6552 | 4.56 | 4.68 | 4.52 | 4.66 | 4.55 | 4.2619 | 4.87 | 4.7021 | > 2 | 4.9074 | 4.6548 | 4.69 | 4.65 | > 1 | > 1 | 4.89 | > 1 | 4.4014 | 4.561 |  | 1
| 1 | 2 | 4.39 | 4.551 | 4.59 | 4.5061 | 4.61 | 4.52 | 4.66 | 4.56 | 4.7 | 4.58 | 4.7 | 4.51 | 4.2563 | 4.86 | 4.4484 | > 2 | 4.9484 | 4.6815 | 4.67 | 4.61 | > 1 | > 1 | 4.45 | > 1 | 4.3643 | 4.611 |  | 39
| 1 | 3 | 4.45 | 4.509 | 4.53 | 4.8313 | 4.65 | 4.49 | 4.6523 | 4.54 | 4.62 | 4.54 | 4.68 | 4.56 | 4.2863 | 5.11 | 4.3231 | > 2 | 4.8733 | 4.7124 | 4.64 | 4.67 | > 1 | > 1 | 4.61 | > 1 | 4.3364 | 4.425 |  | 16
| 1 | 4 | 4.41 | 4.58 | 4.58 | 4.7756 | 4.63 | 4.54 | 4.5587 | 4.56 | 4.55 | 4.51 | 4.69 | 4.53 | 4.2607 | 4.82 | 4.4186 | > 2 | 4.869 | 4.6679 | 4.65 | 4.63 | > 1 | > 1 | 4.53 | > 1 | 4.424 | 4.416 |  | 4.5716
| 1 | 5 | 4.38 | 4.566 | 4.58 | 4.807 | 4.65 | 4.44 | 4.5851 | 4.55 | 4.64 | 4.49 | 4.65 | 4.52 | 4.2625 | 4.77 | 4.3961 | > 2 | 4.9015 | 4.6867 | 4.61 | 4.69 | > 1 | > 1 | 4.68 | > 1 | 4.4019 | 4.445 |  | 162
| 1 | 6 | 4.36 | 4.548 | 4.55 | 4.5572 | 4.58 | 4.6 | 4.5873 | 4.55 | 4.6 | 4.56 | 4.69 | 4.54 | 4.2613 | 4.87 | 4.3669 | > 2 | 4.9191 | 4.6783 | 4.64 | 4.62 | > 1 | > 1 | 4.7 | > 1 | 4.3788 | 4.619
| Mean |  | 4.41 | 4.5525 | 4.5633 | 4.6821 | 4.6233 | 4.5217 | 4.6164 | 4.5533 | 4.6317 | 4.5333 | 4.6783 | 4.535 | 4.2648 | 4.8833 | 4.4425 |  | 4.9031 | 4.6803 | 4.65 | 4.645 |  |  | 4.6433 |  | 4.3845 | 4.5128
| Median |  | 4.4 | 4.556 | 4.565 | 4.6955 | 4.625 | 4.53 | 4.6198 | 4.555 | 4.63 | 4.53 | 4.685 | 4.535 | 4.2616 | 4.865 | 4.4073 |  | 4.9044 | 4.6799 | 4.645 | 4.64 |  |  | 4.645 |  | 4.3901 | 4.503
| Std Dev. |  | 0.042426 | 0.024189 | 0.023381 | 0.13973 | 0.026583 | 0.053821 | 0.044385 | 0.008165 | 0.054559 | 0.033267 | 0.019408 | 0.018708 | 0.010745 | 0.11759 | 0.13426 |  | 0.029605 | 0.019395 | 0.027568 | 0.030822 |  |  | 0.1528 |  | 0.031326 | 0.094785
| Rel.Std.Dev. |  | 0.9621% | 0.5313% | 0.5124% | 2.984% | 0.575% | 1.19% | 0.9615% | 0.1793% | 1.178% | 0.7338% | 0.4148% | 0.4125% | 0.2519% | 2.408% | 3.022% |  | 0.6038% | 0.4144% | 0.5929% | 0.6636% |  |  | 3.291% |  | 0.7145% | 2.1%
| PDM3 |  | -3.534% | -0.4168% | -0.1798% | 2.418% | 1.133% | -1.091% | 0.9819% | -0.3985% | 1.315% | -0.836% | 2.336% | -0.7996% | -6.709% | 6.82% | -2.822% |  | 7.253% | 2.378% | 1.716% | 1.607% |  |  | 1.57% |  | -4.092% | -1.284%
| Z-Score (Absolute) |  | 1.89 | 0.39 | 0.28 | 0.97 | 0.35 | 0.72 | 0.28 | 0.38 | 0.44 | 0.59 | 0.93 | 0.58 | 3.42 | 3.08 | 1.55 | NA | 3.29 | 0.95 | 0.63 | 0.58 | NA | NA | 0.56 | NA | 2.16 | 0.81
| Analytical results for Dy in OREAS 555b (Certified Value 1 ppm) | [28 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 2
| 1 | 1 | 1.016 | 1 | 0.89 | 0.66777 | 0.93072 | 1.17 | [21 empty] | 1
| 1 | 2 | 1.018 | 1 | 0.91 | 0.66164 | 0.93803 | 1.2 | [21 empty] | 10
| 1 | 3 | 0.952 | 1 | 0.96 | 0.7114 | 0.92764 | 1.14 | [21 empty] | 16
| 1 | 4 | 0.981 | 1 | 0.92 | 0.68677 | 0.88187 | 1.23 | [21 empty] | 0.99577
| 1 | 5 | 0.932 | 0.9 | 0.9 | 0.65272 | 0.91654 | 1.15 | [21 empty] | 163
| 1 | 6 | 1.041 | 0.9 | 0.92 | 0.64953 | 0.92301 | 1.18
| Mean |  | 0.99 | 0.96667 | 0.91667 | 0.67164 | 0.91963 | 1.1783
| Median |  | 0.9985 | 1 | 0.915 | 0.66471 | 0.92533 | 1.175
| Std Dev. |  | 0.042308 | 0.05164 | 0.024221 | 0.023543 | 0.019862 | 0.033116
| Rel.Std.Dev. |  | 4.274% | 5.342% | 2.642% | 3.505% | 2.16% | 2.81%
| PDM3 |  | -0.5796% | -2.923% | -7.944% | -32.55% | -7.646% | 18.33%
| Z-Score (Absolute) |  | 0.86 | 0.43 | 0.49 | 4.99 | 0.43 | 4.33
| Analytical results for Er in OREAS 555b (Certified Value 0.41 ppm) | [28 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 2
| 1 | 1 | 0.423 | 0.5 | 0.39 | 0.27798 | 0.40792 | 0.53 | [21 empty] | 1
| 1 | 2 | 0.429 | 0.5 | 0.4 | 0.27871 | 0.43718 | 0.52 | [21 empty] | 11
| 1 | 3 | 0.432 | 0.4 | 0.4 | 0.29612 | 0.3985 | 0.52 | [21 empty] | 16
| 1 | 4 | 0.434 | 0.4 | 0.39 | 0.29101 | 0.38014 | 0.54 | [21 empty] | 0.4069
| 1 | 5 | 0.404 | 0.5 | 0.4 | 0.26829 | 0.42269 | 0.52 | [21 empty] | 164
| 1 | 6 | 0.443 | 0.5 | 0.4 | 0.25714 | 0.39617 | 0.52
| Mean |  | 0.4275 | 0.46667 | 0.39667 | 0.27821 | 0.4071 | 0.525
| Median |  | 0.4305 | 0.5 | 0.4 | 0.27835 | 0.40321 | 0.52
| Std Dev. |  | 0.013248 | 0.05164 | 0.005164 | 0.014334 | 0.02034 | 0.0083666
| Rel.Std.Dev. |  | 3.099% | 11.07% | 1.302% | 5.152% | 4.996% | 1.594%
| PDM3 |  | 5.064% | 14.69% | -2.514% | -31.63% | 0.05033% | 29.03%
| Z-Score (Absolute) |  | 0.67 | NA | 0.34 | 4.26 | 0 | 3.9
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 555b (Certified Value 0.39 ppm) | [28 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 3
| 1 | 1 | 0.415 | 0.4 | 0.34 | 0.20206 | 0.36767 | 0.44 | [21 empty] | 1
| 1 | 2 | 0.378 | 0.4 | 0.35 | 0.19561 | 0.3684 | 0.45 | [21 empty] | 12
| 1 | 3 | 0.378 | 0.4 | 0.36 | 0.20723 | 0.35743 | 0.43 | [21 empty] | 16
| 1 | 4 | 0.385 | 0.4 | 0.34 | 0.20071 | 0.34252 | 0.44 | [21 empty] | 0.38641
| 1 | 5 | 0.331 | 0.4 | 0.34 | 0.19389 | 0.37613 | 0.44 | [21 empty] | 165
| 1 | 6 | 0.411 | 0.4 | 0.35 | 0.18895 | 0.36223 | 0.47
| Mean |  | 0.383 | 0.4 | 0.34667 | 0.19807 | 0.3624 | 0.445
| Median |  | 0.3815 | 0.4 | 0.345 | 0.19816 | 0.36495 | 0.44
| Std Dev. |  | 0.030219 | 6.0809e-17 | 0.008165 | 0.0065347 | 0.011593 | 0.013784
| Rel.Std.Dev. |  | 7.89% | 1.52e-14% | 2.355% | 3.299% | 3.199% | 3.098%
| PDM3 |  | -0.8831% | 3.516% | -10.29% | -48.74% | -6.215% | 15.16%
| Z-Score (Absolute) |  | 0.26 | 0.69 | 0.66 | 4.42 | 0.26 | 1.83
| Analytical results for Fe in OREAS 555b (Certified Value 3.66 wt.%) | [28 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | 3.89 | 3.18 | 3.59 | 3.84 | 3.52 | 3.51 | 3.67 | 3.72 | 3.52 | 3.79 | 3.77 | 3.3123 | 4.155 | 3.4297 | 3.94 | 3.628 | 3.5696 | 3.58 | 3.71 | 3.41 | 3.92 | 3.56 | 3.62 | 3.69 | 3.93 |  |  | 1
| 1 | 2 | 3.81 | 3.48 | 3.5 | 3.84 | 3.47 | 3.57 | 3.72 | 3.75 | 3.55 | 3.73 | 3.75 | 3.397 | 3.875 | 3.504 | 4.01 | 3.7922 | 3.6406 | 3.65 | 3.73 | 3.47 | 4.25 | 3.46 | 3.56 | 3.67 | 3.84 |  |  | #N/A
| 1 | 3 | 3.85 | 3.43 | 3.53 | 3.97 | 3.35 | 3.5 | 3.76 | 3.77 | 3.55 | 3.77 | 3.78 | 3.3874 | 4.183 | 3.7964 | 3.99 | 3.677 | 3.5989 | 3.69 | 3.69 | 3.46 | 4 | 3.43 | 3.63 | 3.65 | 3.82 |  |  | 16
| 1 | 4 | 3.82 | 3.45 | 3.59 | 3.94 | 3.2 | 3.52 | 3.78 | 3.81 | 3.55 | 3.77 | 3.76 | 3.3856 | 3.819 | 3.7154 | 3.95 | 3.6928 | 3.6376 | 3.65 | 3.67 | 3.47 | 4.12 | 3.57 | 3.59 | 3.69 | 3.87 |  |  | 3.6635
| 1 | 5 | 3.82 | 3.61 | 3.63 | 3.74 | 3.32 | 3.43 | 3.81 | 3.79 | 3.59 | 3.76 | 3.72 | 3.3863 | 3.896 | 3.6269 | 3.93 | 3.7555 | 3.5884 | 3.67 | 3.78 | 3.44 | 4.22 | 3.37 | 3.64 | 3.67 | 3.83 |  |  | 166
| 1 | 6 | 3.8 | 3.65 | 3.64 | 3.78 | 3.3 | 3.66 | 3.76 | 3.75 | 3.59 | 3.79 | 3.65 | 3.3764 | 3.763 | 3.8831 | 3.93 | 3.7003 | 3.629 | 3.59 | 3.69 | 3.55 | 4.36 | 3.54 | 3.68 | 3.71 | 3.86
| Mean |  | 3.8317 | 3.4667 | 3.58 | 3.8517 | 3.36 | 3.5317 | 3.75 | 3.765 | 3.5583 | 3.7683 | 3.7383 | 3.3742 | 3.9485 | 3.6592 | 3.9583 | 3.7076 | 3.6107 | 3.6383 | 3.7117 | 3.4667 | 4.145 | 3.4883 | 3.62 | 3.68 | 3.8583
| Median |  | 3.82 | 3.465 | 3.59 | 3.84 | 3.335 | 3.515 | 3.76 | 3.76 | 3.55 | 3.77 | 3.755 | 3.3859 | 3.8855 | 3.6712 | 3.945 | 3.6966 | 3.6139 | 3.65 | 3.7 | 3.465 | 4.17 | 3.5 | 3.625 | 3.68 | 3.85
| Std Dev. |  | 0.033116 | 0.16621 | 0.055136 | 0.089088 | 0.11713 | 0.077309 | 0.04899 | 0.032094 | 0.027142 | 0.022286 | 0.047924 | 0.031006 | 0.17718 | 0.17314 | 0.033714 | 0.058292 | 0.029251 | 0.044008 | 0.0392 | 0.046762 | 0.16441 | 0.08085 | 0.041473 | 0.020976 | 0.039707
| Rel.Std.Dev. |  | 0.8643% | 4.795% | 1.54% | 2.313% | 3.486% | 2.189% | 1.306% | 0.8524% | 0.7628% | 0.5914% | 1.282% | 0.9189% | 4.487% | 4.732% | 0.8517% | 1.572% | 0.8101% | 1.21% | 1.056% | 1.349% | 3.966% | 2.318% | 1.146% | 0.57% | 1.029%
| PDM3 |  | 4.591% | -5.372% | -2.279% | 5.137% | -8.284% | -3.598% | 2.362% | 2.771% | -2.87% | 2.862% | 2.043% | -7.897% | 7.78% | -0.1158% | 8.049% | 1.205% | -1.441% | -0.6863% | 1.315% | -5.372% | 13.14% | -4.781% | -1.187% | 0.451% | 5.319%
| Z-Score (Absolute) |  | 0.84 | 1.18 | 0.55 | 0.95 | 1.77 | 0.82 | 0.39 | 0.47 | 0.67 | 0.49 | 0.32 | 1.7 | 1.49 | 0.12 | 1.54 | 0.15 | 0.38 | 0.23 | 0.18 | 1.18 | 2.58 | 1.06 | 0.33 | 0 | 0.99
| Analytical results for Ga in OREAS 555b (Certified Value 8.22 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 2
| 1 | 1 | 6.3 | 7.8 | 9.86 | 8.92 | 10.4 | 6.86 | 7 | 8.49 | 7.79 | 7.69 | 7.54 | 7.88 | 6.21 | 9.6229 | 6.1336 | 13.8 | 7.8688 | 6.89 | < 10 | 10 | 12.3 | 14.41 | 8 | 9.5 |  |  |  | 1
| 1 | 2 | < 5 | 7.9 | 9.62 | 9.1 | 10.3 | 7.05 | 7.17 | 8.81 | 7.63 | 7.73 | 7.52 | 7.61 | 6.38 | 9.7742 | 6.1001 | 13.9 | 8.1899 | 6.95 | < 10 | 10 | 13.4 | 14.33 | 8 | 9.11 |  |  |  | 13
| 1 | 3 | < 5 | 9 | 9.93 | 9.05 | 10.1 | 6.92 | 7.29 | 8.88 | 7.39 | 7.85 | 7.49 | 7.85 | 6.75 | 10.4346 | 6.2182 | 13.7 | 8.1191 | 7 | < 10 | 10 | 12.9 | 14.1 | 8 | 9.2 |  |  |  | 16
| 1 | 4 | 5.1 | 9.1 | 9.8 | 9.22 | 10.1 | 7.04 | 7.26 | 8.49 | 7.9 | 7.95 | 7.55 | 7.82 | 6.37 | 9.5717 | 6.4143 | 13.9 | 8.2199 | 7 | < 10 | 10 | 12.8 | 13.89 | 8 | 8.9 |  |  |  | 8.2218
| 1 | 5 | < 5 | 10 | 9.86 | 9.04 | 10.4 | 7.14 | 7.22 | 8.79 | 8.13 | 7.77 | 7.28 | 7.77 | 6.34 | 11.0235 | 6.0638 | 13.9 | 8.124 | 6.98 | < 10 | 10 | 9.96 | 13.81 | 8 | 9.33 |  |  |  | 167
| 1 | 6 | 6.5 | 9.2 | 9.84 | 9.29 | 10.1 | 7.36 | 7.19 | 8.91 | 7.73 | 7.81 | 7.27 | 7.79 | 6.44 | 10.6538 | 5.9706 | 13.7 | 8.0064 | 6.75 | < 10 | 10 | 9.66 | 13.7 | 8 | 9.38
| Mean |  | 5.9667 | 8.8333 | 9.8183 | 9.1033 | 10.2333 | 7.0617 | 7.1883 | 8.7283 | 7.7617 | 7.8 | 7.4417 | 7.7867 | 6.415 | 10.1801 | 6.1501 | 13.8167 | 8.088 | 6.9283 |  | 10 | 11.8367 | 14.04 | 8 | 9.2367
| Median |  | 6.3 | 9.05 | 9.85 | 9.075 | 10.2 | 7.045 | 7.205 | 8.8 | 7.76 | 7.79 | 7.505 | 7.805 | 6.375 | 10.1044 | 6.1168 | 13.85 | 8.1215 | 6.965 |  | 10 | 12.55 | 13.995 | 8 | 9.265
| Std Dev. |  | 0.75719 | 0.84063 | 0.10591 | 0.13337 | 0.15055 | 0.17691 | 0.10226 | 0.18978 | 0.24983 | 0.092736 | 0.13075 | 0.095219 | 0.18097 | 0.60761 | 0.15291 | 0.098319 | 0.13018 | 0.096626 |  | 0 | 1.611 | 0.28831 | 0 | 0.21417
| Rel.Std.Dev. |  | 12.69% | 9.517% | 1.079% | 1.465% | 1.471% | 2.505% | 1.423% | 2.174% | 3.219% | 1.189% | 1.757% | 1.223% | 2.821% | 5.969% | 2.486% | 0.7116% | 1.61% | 1.395% |  | 0% | 13.61% | 2.053% | 0% | 2.319%
| PDM3 |  | -27.43% | 7.438% | 19.42% | 10.72% | 24.47% | -14.11% | -12.57% | 6.161% | -5.597% | -5.13% | -9.489% | -5.292% | -21.98% | 23.82% | -25.2% | 68.05% | -1.627% | -15.73% |  | 21.63% | 43.97% | 70.77% | -2.698% | 12.34%
| Z-Score (Absolute) |  | 2.04 | 0.49 | 1.03 | 0.64 | 1.26 | 0.49 | 0.42 | 0.43 | 0.1 | 0.08 | 0.28 | 0.09 | 0.84 | 1.23 | 0.99 | 3.23 | 0.08 | 0.56 | 1.62 | NA | 2.14 | 3.35 | NA | 0.71
Results from laboratories 22 and 27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 555b (Certified Value 2.03 ppm) | [28 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 3
| 1 | 1 | 1.817 | 2.2 | 1.48 | 1.9619 | 2.0897 | 2.17 | [21 empty] | 1
| 1 | 2 | 1.822 | 2.3 | 1.55 | 1.9294 | 2.1214 | 2.15 | [21 empty] | 14
| 1 | 3 | 1.833 | 2 | 1.62 | 2.0214 | 2.1265 | 2.12 | [21 empty] | 16
| 1 | 4 | 1.879 | 2 | 1.51 | 2.0182 | 1.9854 | 2.16 | [21 empty] | 2.025
| 1 | 5 | 1.71 | 2 | 1.54 | 1.9262 | 2.0317 | 2.21 | [21 empty] | 168
| 1 | 6 | 1.891 | 2.1 | 1.54 | 1.9014 | 2.0855 | 2.19
| Mean |  | 1.8253 | 2.1 | 1.54 | 1.9598 | 2.0734 | 2.1667
| Median |  | 1.8275 | 2.05 | 1.54 | 1.9457 | 2.0876 | 2.165
| Std Dev. |  | 0.064239 | 0.12649 | 0.046904 | 0.050345 | 0.054822 | 0.031411
| Rel.Std.Dev. |  | 3.519% | 6.023% | 3.046% | 2.569% | 2.644% | 1.45%
| PDM3 |  | -9.861% | 3.702% | -23.95% | -3.223% | 2.387% | 6.995%
| Z-Score (Absolute) |  | 1.1 | 0.48 | 2.75 | 0.33 | 0.33 | 0.87
| Analytical results for Ge in OREAS 555b (Certified Value 0.095 ppm) | [28 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 18 | 22 | 23 | 25 | 28 | [14 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [14 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 0.1g | 0.5g | [14 empty] | 3
| 1 | 1 | < 0.1 | 0.13 | 0.1 | 0.1 | 0.08 | 0.09 | 0.07 | < 0.01 | < 0.2 | 1.2 | < 0.1 | 0.25 | 0.06 | [14 empty] | 1
| 1 | 2 | < 0.1 | 0.14 | 0.11 | 0.1 | 0.07 | 0.11 | 0.08 | < 0.01 | < 0.2 | 1.1 | < 0.1 | 0.27 | < 0.05 | [14 empty] | 15
| 1 | 3 | < 0.1 | 0.15 | 0.1 | 0.09 | 0.08 | 0.11 | 0.08 | < 0.01 | < 0.2 | 1.1 | < 0.1 | 0.25 | 0.07 | [14 empty] | 16
| 1 | 4 | < 0.1 | 0.15 | 0.1 | 0.1 | 0.08 | 0.12 | 0.07 | < 0.01 | < 0.2 | 1.1 | < 0.1 | 0.24 | < 0.05 | [14 empty] | 0.095476
| 1 | 5 | < 0.1 | 0.16 | 0.11 | 0.09 | 0.08 | 0.1 | 0.08 | < 0.01 | < 0.2 | 1.1 | < 0.1 | 0.1 | 0.06 | [14 empty] | 169
| 1 | 6 | < 0.1 | 0.16 | 0.1 | 0.09 | 0.08 | 0.12 | 0.07 | < 0.01 | < 0.2 | 1.1 | < 0.1 | 0.13 | 0.05
| Mean |  |  | 0.14833 | 0.10333 | 0.095 | 0.078333 | 0.10833 | 0.075 |  |  | 1.1167 |  | 0.20667 | 0.06
| Median |  |  | 0.15 | 0.1 | 0.095 | 0.08 | 0.11 | 0.075 |  |  | 1.1 |  | 0.245 | 0.06
| Std Dev. |  |  | 0.01169 | 0.005164 | 0.0054772 | 0.0040825 | 0.01169 | 0.0054772 |  |  | 0.040825 |  | 0.072296 | 0.008165
| Rel.Std.Dev. |  |  | 7.881% | 4.997% | 5.766% | 5.212% | 10.79% | 7.303% |  |  | 3.656% |  | 34.98% | 13.61%
| PDM3 |  |  | 55.36% | 8.229% | -0.4988% | -17.96% | 13.47% | -21.45% |  |  | 1070% |  | 116.5% | -37.16%
| Z-Score (Absolute) |  | 0.85 | 1.43 | 0.39 | 0.19 | 0.19 | 0.5 | 0.27 | 1.89 | 0.31 | NA | 0.85 | 2.77 | 0.89
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 555b (Certified Value 0.33 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 22 | 23 | 25 | 26 | 28 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.39 | 0.3 | 0.36 | 0.28 | 0.26 | 0.34 | 0.37 | 0.28 | 0.26 | 0.36462 | 0.34 | 0.32459 | 0.55 | 0.3 | 0.26 | 0.33 | 0.3 | 0.37 | [9 empty] | 1
| 1 | 2 | 0.39 | 0.3 | 0.38 | 0.28 | 0.27 | 0.33 | 0.31 | 0.28 | 0.26 | 0.39145 | 0.33 | 0.34698 | 0.47 | 0.3 | 0.29 | 0.36 | 0.31 | 0.37 | [9 empty] | 19
| 1 | 3 | 0.36 | 0.3 | 0.36 | 0.25 | 0.27 | 0.34 | 0.32 | 0.28 | 0.29 | 0.38253 | 0.32 | 0.32719 | 0.44 | 0.3 | 0.31 | 0.38 | 0.3 | 0.36 | [9 empty] | 16
| 1 | 4 | 0.38 | 0.3 | 0.36 | 0.26 | 0.28 | 0.33 | 0.33 | 0.28 | 0.27 | 0.37333 | 0.33 | 0.30428 | 0.41 | 0.3 | 0.32 | 0.37 | 0.32 | 0.37 | [9 empty] | 0.33039
| 1 | 5 | 0.39 | 0.3 | 0.37 | 0.29 | 0.28 | 0.34 | 0.32 | 0.28 | 0.27 | 0.39294 | 0.33 | 0.33141 | 0.4 | 0.3 | 0.35 | 0.26 | 0.3 | 0.37 | [9 empty] | 170
| 1 | 6 | 0.39 | 0.3 | 0.38 | 0.28 | 0.27 | 0.32 | 0.31 | 0.28 | 0.27 | 0.37621 | 0.33 | 0.33148 | 0.41 | 0.3 | 0.33 | 0.3 | 0.32 | 0.34
| Mean |  | 0.38333 | 0.3 | 0.36833 | 0.27333 | 0.27167 | 0.33333 | 0.32667 | 0.28 | 0.27 | 0.38018 | 0.33 | 0.32765 | 0.44667 | 0.3 | 0.31 | 0.33333 | 0.30833 | 0.36333
| Median |  | 0.39 | 0.3 | 0.365 | 0.28 | 0.27 | 0.335 | 0.32 | 0.28 | 0.27 | 0.37937 | 0.33 | 0.3293 | 0.425 | 0.3 | 0.315 | 0.345 | 0.305 | 0.37
| Std Dev. |  | 0.012111 | 0 | 0.0098319 | 0.015055 | 0.0075277 | 0.008165 | 0.022509 | 0 | 0.010954 | 0.010955 | 0.0063246 | 0.013845 | 0.056804 | 0 | 0.031623 | 0.046332 | 0.0098319 | 0.012111
| Rel.Std.Dev. |  | 3.159% | 0% | 2.669% | 5.508% | 2.771% | 2.449% | 6.891% | 0% | 4.057% | 2.882% | 1.917% | 4.225% | 12.72% | 0% | 10.2% | 13.9% | 3.189% | 3.333%
| PDM3 |  | 16.03% | -9.197% | 11.49% | -17.27% | -17.77% | 0.8923% | -1.126% | -15.25% | -18.28% | 15.07% | -0.1166% | -0.8267% | 35.2% | -9.197% | -6.17% | 0.8923% | -6.675% | 9.973%
| Z-Score (Absolute) |  | 0.99 | NA | 0.72 | 1.01 | 1.04 | 0.08 | 0.04 | 0.89 | 1.07 | 0.94 | 0.02 | 0.02 | 2.15 | NA | 0.34 | 0.08 | 0.37 | 0.63
Results from laboratories 5 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 555b (Certified Value 0.06 ppm) | [28 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 18 | 19 | 21 | 22 | 25 | 26 | 27 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 3
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | [6 empty] | 3
| 1 | 1 | < 5 | < 0.005 | < 0.1 | 0.06 | 0.07 | < 0.08 | 0.05 | 0.06 | 0.08 | 0.06 | 0.06 | 0.06 | < 0.05 | < 1 | 0.05 | 1 | 0.751 | 0.04 | 0.08 | 0.06 | 0.057 | [6 empty] | 1
| 1 | 2 | < 5 | 0.062 | < 0.1 | 0.05 | 0.08 | < 0.08 | 0.07 | 0.06 | 0.07 | 0.07 | 0.06 | 0.07 | < 0.05 | < 1 | 0.05 | 1 | 0.759 | 0.04 | 0.09 | < 0.05 | 0.047 | [6 empty] | 20
| 1 | 3 | < 5 | 0.023 | < 0.1 | 0.05 | 0.08 | < 0.08 | 0.06 | 0.06 | 0.07 | 0.05 | 0.05 | 0.06 | < 0.05 | < 1 | 0.06 | 2 | 0.725 | 0.03 | 0.09 | < 0.05 | 0.065 | [6 empty] | 16
| 1 | 4 | < 5 | 0.03 | < 0.1 | 0.06 | 0.07 | < 0.08 | 0.05 | 0.06 | 0.09 | 0.05 | 0.07 | 0.06 | < 0.05 | < 1 | 0.06 | 1 | 0.71 | 0.05 | 0.09 | 0.06 | 0.061 | [6 empty] | 0.060264
| 1 | 5 | < 5 | 0.025 | < 0.1 | 0.05 | 0.08 | < 0.08 | 0.06 | 0.06 | 0.08 | 0.06 | 0.06 | 0.06 | < 0.05 | < 1 | 0.05 | 1 | 0.724 | 0.03 | 0.1 | 0.07 | 0.065 | [6 empty] | 171
| 1 | 6 | < 5 | 0.034 | < 0.1 | 0.07 | 0.08 | < 0.08 | 0.04 | 0.07 | 0.06 | 0.07 | 0.07 | 0.06 | < 0.05 | < 1 | 0.06 | 1 | 0.721 | 0.04 | 0.1 | < 0.05 | 0.054
| Mean |  |  | 0.0348 |  | 0.056667 | 0.076667 |  | 0.055 | 0.061667 | 0.075 | 0.06 | 0.061667 | 0.061667 |  |  | 0.055 | 1.1667 | 0.73167 | 0.038333 | 0.091667 | 0.063333 | 0.058167
| Median |  |  | 0.03 |  | 0.055 | 0.08 |  | 0.055 | 0.06 | 0.075 | 0.06 | 0.06 | 0.06 |  |  | 0.055 | 1 | 0.7245 | 0.04 | 0.09 | 0.06 | 0.059
| Std Dev. |  |  | 0.015802 |  | 0.008165 | 0.005164 |  | 0.010488 | 0.0040825 | 0.010488 | 0.0089443 | 0.0075277 | 0.0040825 |  |  | 0.0054772 | 0.40825 | 0.019012 | 0.0075277 | 0.0075277 | 0.0057735 | 0.0069976
| Rel.Std.Dev. |  |  | 45.41% |  | 14.41% | 6.736% |  | 19.07% | 6.62% | 13.98% | 14.91% | 12.21% | 6.62% |  |  | 9.959% | 34.99% | 2.598% | 19.64% | 8.212% | 9.116% | 12.03%
| PDM3 |  |  | -42.25% |  | -5.969% | 27.22% |  | -8.735% | 2.328% | 24.45% | -0.4379% | 2.328% | 2.328% |  |  | -8.735% | 1836% | 1114% | -36.39% | 52.11% | 5.093% | -3.48%
| Z-Score (Absolute) |  | 106.76 | 1.3 | 0.4 | 0.11 | 0.77 | 0.83 | 0.18 | 0.11 | 0.7 | 0.04 | 0.11 | 0.11 | 1.49 | 19.29 | 0.18 | NA | 29.42 | 0.91 | 1.43 | 0.65 | 0.04
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 555b (Certified Value 0.17 ppm) | [28 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 04 | 05 | 14 | 18 | 26 | [22 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [22 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.2g | 01g | [22 empty] | 2
| 1 | 1 | 0.164 | 0.2 | 0.12 | 0.14839 | 0.19 | [22 empty] | 1
| 1 | 2 | 0.166 | 0.2 | 0.13 | 0.15525 | 0.21 | [22 empty] | 3
| 1 | 3 | 0.166 | 0.2 | 0.13 | 0.15064 | 0.19 | [22 empty] | 16
| 1 | 4 | 0.17 | 0.2 | 0.13 | 0.14232 | 0.19 | [22 empty] | 0.16772
| 1 | 5 | 0.168 | 0.2 | 0.13 | 0.15172 | 0.19 | [22 empty] | 172
| 1 | 6 | 0.17 | 0.2 | 0.13 | 0.14913 | 0.19
| Mean |  | 0.16733 | 0.2 | 0.12833 | 0.14958 | 0.19333
| Median |  | 0.167 | 0.2 | 0.13 | 0.14989 | 0.19
| Std Dev. |  | 0.0024221 | 3.0405e-17 | 0.0040825 | 0.0042934 | 0.008165
| Rel.Std.Dev. |  | 1.447% | 1.52e-14% | 3.181% | 2.87% | 4.223%
| PDM3 |  | -0.2277% | 19.25% | -23.48% | -10.82% | 15.27%
| Z-Score (Absolute) |  | 0 | 0.85 | 1.01 | 0.46 | 0.67
| Analytical results for In in OREAS 555b (Certified Value 1.83 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 20 | 22 | 23 | 25 | 26 | 28 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [9 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [9 empty] | 3
| 1 | 1 | < 5 | 1.9 | 1.72 | 1.71 | 1.7938 | 1.83 | 1.86 | 1.875 | 1.825 | 1.85 | 1.6778 | 2.09 | 2 | 1.73 | 1.59 | 2 | 2.09 | 1.902 | [9 empty] | 1
| 1 | 2 | < 5 | 1.9 | 1.74 | 1.74 | 1.7854 | 1.85 | 1.84 | 1.88 | 1.825 | 1.85 | 1.6838 | 2.13 | 2 | 1.7 | 1.679 | 2 | 2.09 | 1.832 | [9 empty] | 4
| 1 | 3 | < 5 | 1.87 | 1.69 | 1.73 | 1.7984 | 1.835 | 1.855 | 1.825 | 1.83 | 1.85 | 1.8506 | 2.1 | 2 | 1.69 | 1.577 | 1.95 | 2.02 | 1.836 | [9 empty] | 16
| 1 | 4 | < 5 | 1.93 | 1.62 | 1.78 | 1.8062 | 1.865 | 1.87 | 1.94 | 1.84 | 1.845 | 1.7682 | 2.12 | 2 | 1.76 | 1.625 | 1.98 | 2.04 | 1.814 | [9 empty] | 1.827
| 1 | 5 | < 5 | 1.79 | 1.68 | 1.78 | 1.7877 | 1.845 | 1.875 | 2.02 | 1.835 | 1.805 | 1.7723 | 2.09 | 2 | 1.71 | 1.723 | 1.99 | 2.04 | 1.847 | [9 empty] | 173
| 1 | 6 | < 5 | 1.9 | 1.68 | 1.82 | 1.8159 | 1.85 | 1.865 | 1.865 | 1.825 | 1.8 | 1.8714 | 2.08 | 2 | 1.74 | 1.709 | 1.97 | 2.07 | 1.812
| Mean |  |  | 1.8817 | 1.6883 | 1.76 | 1.7979 | 1.8458 | 1.8608 | 1.9008 | 1.83 | 1.8333 | 1.7707 | 2.1017 | 2 | 1.7217 | 1.6505 | 1.9817 | 2.0583 | 1.8405
| Median |  |  | 1.9 | 1.685 | 1.76 | 1.7961 | 1.8475 | 1.8625 | 1.8775 | 1.8275 | 1.8475 | 1.7703 | 2.095 | 2 | 1.72 | 1.652 | 1.985 | 2.055 | 1.834
| Std Dev. |  |  | 0.048751 | 0.041191 | 0.040497 | 0.01157 | 0.012416 | 0.012416 | 0.069095 | 0.0063246 | 0.024014 | 0.080892 | 0.019408 | 0 | 0.026394 | 0.061973 | 0.019408 | 0.029269 | 0.032959
| Rel.Std.Dev. |  |  | 2.591% | 2.44% | 2.301% | 0.6435% | 0.6727% | 0.6672% | 3.635% | 0.3456% | 1.31% | 4.568% | 0.9235% | 0% | 1.533% | 3.755% | 0.9794% | 1.422% | 1.791%
| PDM3 |  |  | 2.995% | -7.587% | -3.665% | -1.59% | 1.033% | 1.854% | 4.044% | 0.1668% | 0.3492% | -3.079% | 15.04% | 9.472% | -5.763% | -9.658% | 8.468% | 12.66% | 0.7415%
| Z-Score (Absolute) |  | 6.37 | 0.4 | 1.47 | 0.78 | 0.41 | 0.05 | 0.2 | 0.58 | 0.1 | 0.07 | 0.67 | 2.52 | NA | 1.15 | 1.84 | 1.36 | 2.1 | 0
Results from laboratory 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ir in OREAS 555b (Indicative Value < 0.003 ppm) | [28 empty] | Indicative
Ir | Round | Replicate | Lab | [26 empty] | 1
| No | No | 07 | [26 empty] | ppm
|  |  | AR*MS | [26 empty] | 3
|  |  | 0.5g | [26 empty] | 3
| 1 | 1 | < 0.003 | [26 empty] | 1
| 1 | 2 | < 0.003 | [26 empty] | 24
| 1 | 3 | < 0.003 | [26 empty] | 16
| 1 | 4 | < 0.003 | [26 empty] | < 0.003
| 1 | 5 | < 0.003 | [26 empty] | 30
| 1 | 6 | < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 555b (Certified Value 0.196 wt.%) | [28 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.01 | 0.16 | 0.23 | 0.2061 | 0.26 | 0.16 | 0.17 | 0.21 | 0.18 | 0.19 | 0.19 | 0.2095 | 0.1428 | 0.21969 | 0.307 | 0.19237 | 0.14 | 0.17 | 0.19 | 0.213 | 0.37 | 0.2212 | 0.22 | 0.21 |  |  |  | 1
| 1 | 2 | 0.01 | 0.18 | 0.22 | 0.2087 | 0.25 | 0.16 | 0.18 | 0.22 | 0.18 | 0.18 | 0.2 | 0.2018 | 0.1345 | 0.18321 | 0.32 | 0.1999 | 0.14 | 0.17 | 0.2 | 0.235 | 0.498 | 0.229 | 0.21 | 0.2 |  |  |  | #N/A
| 1 | 3 | 0.01 | 0.19 | 0.22 | 0.2144 | 0.25 | 0.15 | 0.18 | 0.22 | 0.18 | 0.19 | 0.19 | 0.2049 | 0.1428 | 0.20706 | 0.318 | 0.19934 | 0.14 | 0.16 | 0.2 | 0.225 | 0.463 | 0.213 | 0.22 | 0.2 |  |  |  | 16
| 1 | 4 | 0.01 | 0.2 | 0.23 | 0.2075 | 0.25 | 0.15 | 0.18 | 0.22 | 0.18 | 0.19 | 0.2 | 0.2057 | 0.1303 | 0.18872 | 0.303 | 0.19584 | 0.14 | 0.16 | 0.2 | 0.233 | 0.49 | 0.2377 | 0.22 | 0.2 |  |  |  | 0.19566
| 1 | 5 | 0.01 | 0.21 | 0.24 | 0.2088 | 0.25 | 0.15 | 0.18 | 0.22 | 0.18 | 0.18 | 0.19 | 0.2069 | 0.1303 | 0.22058 | 0.321 | 0.19137 | 0.14 | 0.17 | 0.19 | 0.238 | 0.288 | 0.219 | 0.22 | 0.2 |  |  |  | 174
| 1 | 6 | 0.01 | 0.21 | 0.24 | 0.2054 | 0.23 | 0.16 | 0.18 | 0.21 | 0.18 | 0.19 | 0.19 | 0.2072 | 0.1254 | 0.21758 | 0.326 | 0.19244 | 0.14 | 0.17 | 0.19 | 0.244 | 0.284 | 0.2245 | 0.22 | 0.21
| Mean |  | 0.01 | 0.19167 | 0.23 | 0.20848 | 0.24833 | 0.155 | 0.17833 | 0.21667 | 0.18 | 0.18667 | 0.19333 | 0.206 | 0.13435 | 0.20614 | 0.31583 | 0.19521 | 0.14 | 0.16667 | 0.195 | 0.23133 | 0.39883 | 0.22407 | 0.21833 | 0.20333
| Median |  | 0.01 | 0.195 | 0.23 | 0.2081 | 0.25 | 0.155 | 0.18 | 0.22 | 0.18 | 0.19 | 0.19 | 0.2063 | 0.1324 | 0.21232 | 0.319 | 0.19414 | 0.14 | 0.17 | 0.195 | 0.234 | 0.4165 | 0.22285 | 0.22 | 0.2
| Std Dev. |  | 0 | 0.019408 | 0.0089443 | 0.0032034 | 0.0098319 | 0.0054772 | 0.0040825 | 0.005164 | 3.0405e-17 | 0.005164 | 0.005164 | 0.0025861 | 0.0071517 | 0.01645 | 0.0088863 | 0.0037386 | 0 | 0.005164 | 0.0054772 | 0.01093 | 0.098554 | 0.0085645 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 0% | 10.13% | 3.889% | 1.537% | 3.959% | 3.534% | 2.289% | 2.383% | 1.689e-14% | 2.766% | 2.671% | 1.255% | 5.323% | 7.98% | 2.814% | 1.915% | 0% | 3.098% | 2.809% | 4.725% | 24.71% | 3.822% | 1.87% | 2.54%
| PDM3 |  | -94.89% | -2.042% | 17.55% | 6.552% | 26.92% | -20.78% | -8.857% | 10.73% | -8.005% | -4.598% | -1.19% | 5.283% | -31.34% | 5.356% | 61.42% | -0.2309% | -28.45% | -14.82% | -0.3387% | 18.23% | 103.8% | 14.52% | 11.59% | 3.92%
| Z-Score (Absolute) |  | 6.35 | 0.26 | 1.03 | 0.31 | 1.65 | 1.48 | 0.7 | 0.58 | 0.65 | 0.42 | 0.2 | 0.23 | 2.18 | 0.23 | 3.91 | 0.14 | 1.99 | 1.09 | 0.14 | 1.08 | 6.69 | 0.83 | 0.64 | 0.14
| Analytical results for La in OREAS 555b (Certified Value 20.4 ppm) | [28 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 1
| 1 | 1 | 28.4 | 17.9 | 21.6 | 21.813 | 26.4 | 15.6 | 18.5 | 21.6 | 19 | 18.8 | 19.7 | 18.14 | 16 | 21.5028 | 10.8041 | 25.4 | 18.7829 | 15 | 18 | 18.1 | 34.7 | 25.24 | 20.2 | 21 |  |  |  | 1
| 1 | 2 | 24.4 | 19.9 | 21.2 | 21.669 | 25.2 | 16.1 | 18.8 | 22.6 | 18.9 | 19.1 | 20 | 17.35 | 14 | 20.1905 | 10.6765 | 25.1 | 19.2469 | 15.1 | 18 | 19.6 | 38.9 | 26.14 | 19.8 | 20.6 |  |  |  | 5
| 1 | 3 | 23.5 | 20.9 | 21.8 | 21.005 | 25.3 | 15.8 | 19.1 | 22.6 | 18.3 | 19.1 | 19.3 | 17.84 | 16 | 22.1753 | 10.6928 | 26.2 | 19.5463 | 15.1 | 18 | 19.7 | 38.3 | 24.91 | 19.8 | 20.8 |  |  |  | 16
| 1 | 4 | 24.3 | 21.3 | 21.5 | 21.745 | 23.4 | 16.3 | 19.5 | 22.1 | 18.6 | 19.4 | 20.3 | 17.96 | 14 | 21.5681 | 11.0746 | 24.7 | 18.3712 | 15 | 19 | 20.3 | 37.1 | 25.6 | 20.5 | 20.2 |  |  |  | 20.3761
| 1 | 5 | 24.4 | 22.6 | 22.3 | 20.628 | 25 | 16.3 | 19.8 | 23 | 19.3 | 19.2 | 19.4 | 17.93 | 14 | 24.0038 | 10.3601 | 24.9 | 18.4841 | 15.3 | 18 | 21.1 | 27.5 | 25.31 | 20.6 | 21.2 |  |  |  | 175
| 1 | 6 | 28.6 | 21.9 | 22.4 | 22.142 | 24.2 | 16.7 | 19 | 22.6 | 18.4 | 19.4 | 18.9 | 18.08 | 13 | 23.5405 | 9.9842 | 24.3 | 18.9769 | 15.3 | 18 | 20.4 | 27.4 | 25.74 | 20.5 | 21.1
| Mean |  | 25.6 | 20.75 | 21.8 | 21.5003 | 24.9167 | 16.1333 | 19.1167 | 22.4167 | 18.75 | 19.1667 | 19.6 | 17.8833 | 14.5 | 22.1635 | 10.5987 | 25.1 | 18.9014 | 15.1333 | 18.1667 | 19.8667 | 33.9833 | 25.49 | 20.2333 | 20.8167
| Median |  | 24.4 | 21.1 | 21.7 | 21.707 | 25.1 | 16.2 | 19.05 | 22.6 | 18.75 | 19.15 | 19.55 | 17.945 | 14 | 21.8717 | 10.6846 | 25 | 18.8799 | 15.1 | 18 | 20 | 35.9 | 25.455 | 20.35 | 20.9
| Std Dev. |  | 2.2724 | 1.6682 | 0.46904 | 0.56641 | 1.0245 | 0.39328 | 0.47081 | 0.4916 | 0.38341 | 0.22509 | 0.50596 | 0.28247 | 1.2247 | 1.4119 | 0.37897 | 0.65422 | 0.44952 | 0.13663 | 0.40825 | 1.0211 | 5.2613 | 0.43091 | 0.36148 | 0.37103
| Rel.Std.Dev. |  | 8.877% | 8.04% | 2.152% | 2.634% | 4.112% | 2.438% | 2.463% | 2.193% | 2.045% | 1.174% | 2.581% | 1.579% | 8.447% | 6.37% | 3.576% | 2.606% | 2.378% | 0.9028% | 2.247% | 5.14% | 15.48% | 1.69% | 1.787% | 1.782%
| PDM3 |  | 25.64% | 1.835% | 6.988% | 5.517% | 22.28% | -20.82% | -6.181% | 10.01% | -7.981% | -5.936% | -3.809% | -12.23% | -28.84% | 8.772% | -47.98% | 23.18% | -7.238% | -25.73% | -10.84% | -2.5% | 66.78% | 25.1% | -0.7009% | 2.162%
| Z-Score (Absolute) |  | 1.87 | 0.24 | 0.59 | 0.49 | 1.64 | 1.32 | 0.31 | 0.8 | 0.44 | 0.3 | 0.15 | 0.73 | 1.87 | 0.71 | 3.19 | 1.7 | 0.39 | 1.66 | 0.64 | 0.06 | 4.7 | 1.84 | 0.06 | 0.26
| Analytical results for Li in OREAS 555b (Certified Value 35.3 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | [6 empty] | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [6 empty] | 1
| 1 | 1 | 44 | 34.6 | 43.5 | 19 | 28.3 | 34.7 | 29.7 | 28.4 | 31 | 44.78 | 1 | 31.0273 | 42.4 | 30.9835 | 41.7 | 26 | 31.6 | 37 | 42.3 | 45 | 33.4 | [6 empty] | 1
| 1 | 2 | 43.2 | 34.3 | 45.1 | 20.3 | 28.9 | 35.8 | 29.9 | 28.5 | 30.8 | 44.06 | < 1 | 27.8752 | 43.4 | 32.9247 | 42.6 | 26 | 32.2 | 43 | 45.5 | 44 | 33.7 | [6 empty] | 33
| 1 | 3 | 42.8 | 34.5 | 43.5 | 18.8 | 29.8 | 36 | 29 | 29.2 | 30.9 | 44.26 | 1 | 30.856 | 43.1 | 31.8755 | 43.7 | 25 | 32.3 | 42 | 41.1 | 45 | 33.4 | [6 empty] | 16
| 1 | 4 | 42.8 | 35.8 | 42.5 | 19.7 | 30.1 | 35.8 | 29.3 | 29.2 | 31.3 | 44.38 | < 1 | 28.2024 | 42.4 | 31.4136 | 42.5 | 25 | 31.9 | 43 | 44 | 45 | 30.8 | [6 empty] | 35.3459
| 1 | 5 | 45.1 | 34 | 43.1 | 20 | 30.4 | 35.8 | 29.9 | 28.9 | 30.5 | 44.46 | < 1 | 29.6438 | 42.5 | 32.1068 | 42.8 | 26 | 30.6 | 43 | 36 | 44 | 33.6 | [6 empty] | 176
| 1 | 6 | 45.2 | 35.4 | 43.2 | 20.2 | 29.9 | 35.5 | 29.8 | 29.6 | 29.6 | 44.63 | < 1 | 31.0137 | 41.7 | 31.4326 | 43.5 | 26 | 31.9 | 44 | 36.7 | 46 | 31.8
| Mean |  | 43.85 | 34.7667 | 43.4833 | 19.6667 | 29.5667 | 35.6 | 29.6 | 28.9667 | 30.6833 | 44.4283 | 1 | 29.7697 | 42.5833 | 31.7895 | 42.8 | 25.6667 | 31.75 | 42 | 40.9333 | 44.8333 | 32.7833
| Median |  | 43.6 | 34.55 | 43.35 | 19.85 | 29.85 | 35.8 | 29.75 | 29.05 | 30.85 | 44.42 | 1 | 30.2499 | 42.45 | 31.6541 | 42.7 | 26 | 31.9 | 43 | 41.7 | 45 | 33.4
| Std Dev. |  | 1.0986 | 0.68896 | 0.87274 | 0.6314 | 0.79917 | 0.46904 | 0.36878 | 0.45898 | 0.59133 | 0.25756 | 0 | 1.4402 | 0.59805 | 0.68046 | 0.72664 | 0.5164 | 0.61563 | 2.5298 | 3.8578 | 0.75277 | 1.1974
| Rel.Std.Dev. |  | 2.505% | 1.982% | 2.007% | 3.211% | 2.703% | 1.318% | 1.246% | 1.585% | 1.927% | 0.5797% | 0% | 4.838% | 1.404% | 2.141% | 1.698% | 2.012% | 1.939% | 6.023% | 9.425% | 1.679% | 3.652%
| PDM3 |  | 24.06% | -1.639% | 23.02% | -44.36% | -16.35% | 0.719% | -16.26% | -18.05% | -13.19% | 25.7% | -97.17% | -15.78% | 20.48% | -10.06% | 21.09% | -27.38% | -10.17% | 18.83% | 15.81% | 26.84% | -7.25%
| Z-Score (Absolute) |  | 1.05 | 0.19 | 1.01 | 1.24 | 0.3 | 0.27 | 0.3 | 0.36 | 0.2 | 1.1 | 3.04 | 0.29 | 0.93 | 0.09 | 0.95 | 0.67 | 0.1 | 0.87 | 0.77 | 1.14 | 0
| Analytical results for Lu in OREAS 555b (Indicative Value 0.058 ppm) | [28 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 04 | 05 | 18 | 23 | 26 | [22 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [22 empty] | 3
|  |  | 01g | 0.5g | 0.2g | 0.25g | 01g | [22 empty] | 3
| 1 | 1 | 0.055 | < 0.1 | 0.054368 | 0.05 | 0.07 | [22 empty] | 1
| 1 | 2 | 0.05 | < 0.1 | 0.053276 | 0.06 | 0.06 | [22 empty] | 25
| 1 | 3 | 0.05 | < 0.1 | 0.05116 | 0.06 | 0.06 | [22 empty] | 16
| 1 | 4 | 0.052 | < 0.1 | < 0.05 | 0.06 | 0.07 | [22 empty] | 0.057506
| 1 | 5 | 0.051 | < 0.1 | 0.05334 | 0.07 | 0.06 | [22 empty] | 31
| 1 | 6 | 0.054 | < 0.1 | 0.052977 | 0.07 | 0.06
| Mean |  | 0.052 |  | 0.053024 | 0.061667 | 0.063333
| Median |  | 0.0515 |  | 0.053276 | 0.06 | 0.06
| Std Dev. |  | 0.0020976 |  | 0.0011666 | 0.0075277 | 0.005164
| Rel.Std.Dev. |  | 4.034% |  | 2.2% | 12.21% | 8.154%
| PDM3 |  | -9.575% |  | -7.794% | 7.235% | 10.13%
| Z-Score (Absolute) |  | 0 | 0.37 | 0.67 | 1.79 | 2.1
| Analytical results for Mg in OREAS 555b (Certified Value 2.42 wt.%) | [28 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 2.53 | 2.22 | 2.65 | 2.53 | 2.33 | 1.83 | 2.18 | 2.45 | 2.29 | 2.24 | 2.3 | 2.3258 | 2.056 | 2.7534 | 2.85 | 2.427 | 2.11 | 1.99 | 2.36 | 2.573 | 2.6 | 2.8 | 2.38 | 2.53 |  |  |  | 1
| 1 | 2 | 2.37 | 2.49 | 2.67 | 2.55 | 2.25 | 1.84 | 2.23 | 2.49 | 2.3 | 2.23 | 2.28 | 2.3383 | 1.93 | 2.5184 | 2.92 | 2.5252 | 2.08 | 2 | 2.37 | 2.696 | 2.64 | 2.75 | 2.33 | 2.47 |  |  |  | #N/A
| 1 | 3 | 2.35 | 2.54 | 2.64 | 2.63 | 2.19 | 1.79 | 2.27 | 2.49 | 2.26 | 2.28 | 2.31 | 2.3375 | 2.081 | 2.6666 | 2.92 | 2.4471 | 2.12 | 1.98 | 2.4 | 2.577 | 2.55 | 2.79 | 2.35 | 2.47 |  |  |  | 16
| 1 | 4 | 2.33 | 2.57 | 2.64 | 2.61 | 2.26 | 1.81 | 2.3 | 2.51 | 2.28 | 2.28 | 2.31 | 2.3388 | 1.869 | 2.5197 | 2.85 | 2.4223 | 2.08 | 1.97 | 2.37 | 2.655 | 2.7 | 2.8 | 2.37 | 2.47 |  |  |  | 2.4173
| 1 | 5 | 2.35 | 2.72 | 2.69 | 2.56 | 2.3 | 1.74 | 2.31 | 2.5 | 2.32 | 2.27 | 2.27 | 2.3374 | 1.888 | 2.5615 | 2.83 | 2.4993 | 2.09 | 2.03 | 2.39 | 2.69 | 2.37 | 2.79 | 2.37 | 2.52 |  |  |  | 177
| 1 | 6 | 2.49 | 2.71 | 2.69 | 2.51 | 2.24 | 1.88 | 2.27 | 2.49 | 2.31 | 2.3 | 2.22 | 2.3376 | 1.839 | 2.7609 | 2.82 | 2.4589 | 2.02 | 1.99 | 2.36 | 2.791 | 2.47 | 2.86 | 2.4 | 2.53
| Mean |  | 2.4033 | 2.5417 | 2.6633 | 2.565 | 2.2617 | 1.815 | 2.26 | 2.4883 | 2.2933 | 2.2667 | 2.2817 | 2.3359 | 1.9438 | 2.6301 | 2.865 | 2.4633 | 2.0833 | 1.9933 | 2.375 | 2.6637 | 2.555 | 2.7983 | 2.3667 | 2.4983
| Median |  | 2.36 | 2.555 | 2.66 | 2.555 | 2.255 | 1.82 | 2.27 | 2.49 | 2.295 | 2.275 | 2.29 | 2.3376 | 1.909 | 2.614 | 2.85 | 2.453 | 2.085 | 1.99 | 2.37 | 2.6725 | 2.575 | 2.795 | 2.37 | 2.495
| Std Dev. |  | 0.084538 | 0.1828 | 0.023381 | 0.046368 | 0.048751 | 0.047645 | 0.048166 | 0.020412 | 0.021602 | 0.026583 | 0.034303 | 0.0049775 | 0.10129 | 0.11226 | 0.044159 | 0.040995 | 0.035024 | 0.020656 | 0.016432 | 0.08217 | 0.11979 | 0.035449 | 0.024221 | 0.031252
| Rel.Std.Dev. |  | 3.518% | 7.192% | 0.8779% | 1.808% | 2.156% | 2.625% | 2.131% | 0.8203% | 0.942% | 1.173% | 1.503% | 0.2131% | 5.211% | 4.268% | 1.541% | 1.664% | 1.681% | 1.036% | 0.6919% | 3.085% | 4.689% | 1.267% | 1.023% | 1.251%
| PDM3 |  | -0.5758% | 5.147% | 10.18% | 6.112% | -6.436% | -24.91% | -6.505% | 2.941% | -5.126% | -6.23% | -5.609% | -3.365% | -19.59% | 8.805% | 18.52% | 1.904% | -13.81% | -17.54% | -1.748% | 10.19% | 5.699% | 15.77% | -2.093% | 3.354%
| Z-Score (Absolute) |  | 0.07 | 0.73 | 1.31 | 0.84 | 0.61 | 2.75 | 0.62 | 0.47 | 0.46 | 0.59 | 0.51 | 0.26 | 2.13 | 1.15 | 2.28 | 0.35 | 1.46 | 1.9 | 0.07 | 1.31 | 0.79 | 1.96 | 0.11 | 0.52
| Analytical results for Mn in OREAS 555b (Certified Value 0.089 wt.%) | [28 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.08982 | 0.0827 | 0.0878 | 0.0916 | 0.0903 | 0.082 | 0.0864 | 0.0899 | 0.091 | 0.0868 | 0.0887 | 0.0841 | 0.0906 | 0.091659 | 0.0829 | 0.088786 | 0.0845 | 0.0903 | 0.0714 | 0.1006 | 0.0838 | 0.0904 | 0.0926 | 0.0928 |  |  |  | 1
| 1 | 2 | 0.09134 | 0.0945 | 0.0843 | 0.0937 | 0.0892 | 0.0856 | 0.0882 | 0.0981 | 0.0906 | 0.087 | 0.0889 | 0.0795 | 0.0953 | 0.091704 | 0.0867 | 0.088367 | 0.0863 | 0.088 | 0.0701 | 0.1044 | 0.083 | 0.0889 | 0.0909 | 0.0926 |  |  |  | #N/A
| 1 | 3 | 0.08708 | 0.0922 | 0.0874 | 0.0964 | 0.0832 | 0.0809 | 0.0897 | 0.0958 | 0.0929 | 0.0871 | 0.0897 | 0.0826 | 0.0906 | 0.092838 | 0.0868 | 0.087671 | 0.0882 | 0.0899 | 0.0687 | 0.1017 | 0.0823 | 0.0913 | 0.0909 | 0.0903 |  |  |  | 16
| 1 | 4 | 0.08674 | 0.0893 | 0.0849 | 0.0947 | 0.0833 | 0.0814 | 0.0916 | 0.0941 | 0.0914 | 0.0868 | 0.092 | 0.0833 | 0.0844 | 0.09136 | 0.0834 | 0.08853 | 0.0883 | 0.0866 | 0.0695 | 0.1028 | 0.0834 | 0.0893 | 0.0922 | 0.0919 |  |  |  | 0.088654
| 1 | 5 | 0.09146 | 0.0975 | 0.0872 | 0.0895 | 0.084 | 0.081 | 0.0899 | 0.0937 | 0.0955 | 0.0895 | 0.0908 | 0.0817 | 0.0867 | 0.092632 | 0.0827 | 0.091138 | 0.0828 | 0.0898 | 0.0677 | 0.1028 | 0.0805 | 0.0916 | 0.0918 | 0.0925 |  |  |  | 178
| 1 | 6 | 0.08843 | 0.0928 | 0.0884 | 0.0917 | 0.0834 | 0.0967 | 0.0906 | 0.0916 | 0.0929 | 0.0866 | 0.0927 | 0.0828 | 0.0821 | 0.092382 | 0.0823 | 0.090382 | 0.0855 | 0.0907 | 0.07 | 0.1104 | 0.0848 | 0.0956 | 0.0922 | 0.0914
| Mean |  | 0.089145 | 0.0915 | 0.086667 | 0.092933 | 0.085567 | 0.0846 | 0.0894 | 0.093867 | 0.092383 | 0.0873 | 0.090467 | 0.082333 | 0.088283 | 0.092096 | 0.084133 | 0.089146 | 0.085933 | 0.089217 | 0.069567 | 0.10378 | 0.082967 | 0.091183 | 0.091767 | 0.091917
| Median |  | 0.089125 | 0.0925 | 0.0873 | 0.0927 | 0.0837 | 0.0817 | 0.0898 | 0.0939 | 0.09215 | 0.0869 | 0.09025 | 0.0827 | 0.08865 | 0.092043 | 0.08315 | 0.088658 | 0.0859 | 0.08985 | 0.06975 | 0.1028 | 0.0832 | 0.09085 | 0.092 | 0.0922
| Std Dev. |  | 0.0020588 | 0.0050884 | 0.0016633 | 0.0024825 | 0.0032709 | 0.0061816 | 0.0018472 | 0.0029194 | 0.001806 | 0.0010918 | 0.0016501 | 0.0015983 | 0.0048122 | 0.00060115 | 0.0020578 | 0.0013257 | 0.0021417 | 0.0015817 | 0.0012707 | 0.0034805 | 0.0014679 | 0.0024112 | 0.0007174 | 0.00094534
| Rel.Std.Dev. |  | 2.309% | 5.561% | 1.919% | 2.671% | 3.823% | 7.307% | 2.066% | 3.11% | 1.955% | 1.251% | 1.824% | 1.941% | 5.451% | 0.6527% | 2.446% | 1.487% | 2.492% | 1.773% | 1.827% | 3.354% | 1.769% | 2.644% | 0.7818% | 1.028%
| PDM3 |  | 0.5538% | 3.21% | -2.242% | 4.827% | -3.482% | -4.573% | 0.8415% | 5.88% | 4.207% | -1.527% | 2.045% | -7.13% | -0.4181% | 3.882% | -5.099% | 0.5545% | -3.069% | 0.6347% | -21.53% | 17.07% | -6.415% | 2.853% | 3.511% | 3.68%
| Z-Score (Absolute) |  | 0.01 | 0.55 | 0.6 | 0.9 | 0.86 | 1.09 | 0.05 | 1.12 | 0.76 | 0.45 | 0.31 | 1.63 | 0.21 | 0.7 | 1.2 | 0.01 | 0.78 | 0.01 | 4.68 | 3.49 | 1.48 | 0.48 | 0.62 | 0.65
| Analytical results for Mo in OREAS 555b (Certified Value 16.2 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 2
| 1 | 1 | 4.9 | 15.3 | 16.3 | 16.3 | 15.4 | 16.5 | 16.485 | 16.6 | 16.7 | 15.6 | 16.7 | 15.6 | 14.48 | 16.5 | 16.105 | 16.1 | 17.0455 | 17.1 | 16.2 | 16 | 14.6 | 17.13 | 16.5 | 17.2 | 15.5 | 16.43 |  | 1
| 1 | 2 | 4.2 | 15.06 | 15.8 | 16.2 | 14.6 | 16.9 | 16.5904 | 16.9 | 16.95 | 15.6 | 16.55 | 15.65 | 14.81 | 16.8 | 16.3542 | 16.2 | 17.1786 | 16.5 | 16.63 | 16 | 14.5 | 18.42 | 15.1 | 17.3 | 15.6 | 16.56 |  | 7
| 1 | 3 | 4.3 | 16.15 | 16.2 | 15.9 | 14.4 | 16.5 | 16.5852 | 16.8 | 17.1 | 15.25 | 16.65 | 15.5 | 14.64 | 17.4 | 16.267 | 16.2 | 17.2807 | 17.1 | 16.86 | 16 | 14.7 | 17.41 | 15.2 | 16.8 | 15.2 | 16.26 |  | 16
| 1 | 4 | 5.7 | 15.63 | 16.1 | 16.3 | 13.7 | 17.1 | 16.7794 | 16.95 | 17 | 16.15 | 16.6 | 15.55 | 14.71 | 16.8 | 16.418 | 16.1 | 16.9521 | 17.1 | 16.19 | 16 | 15.1 | 18.17 | 15.2 | 16.9 | 15.7 | 15.93 |  | 16.2231
| 1 | 5 | 5.4 | 16.43 | 16.5 | 15.9 | 14.2 | 17.3 | 16.5061 | 17.3 | 17.1 | 16.4 | 16.55 | 15.1 | 14.76 | 16.6 | 16.1486 | 16.1 | 17.001 | 17.3 | 16.04 | 16 | 15 | 18.85 | 15.7 | 16.8 | 15.5 | 16.28 |  | 179
| 1 | 6 | 5.8 | 15.7 | 16.5 | 16.5 | 14.1 | 17.5 | 16.0366 | 16.85 | 16.8 | 15.6 | 16.55 | 15.25 | 14.72 | 16.8 | 16.0123 | 16.2 | 17.0515 | 17.5 | 16.39 | 16 | 15.1 | 18.6 | 15.9 | 17.4 | 15.7 | 15.53
| Mean |  | 5.05 | 15.7117 | 16.2333 | 16.1833 | 14.4 | 16.9667 | 16.4971 | 16.9 | 16.9417 | 15.7667 | 16.6 | 15.4417 | 14.6867 | 16.8167 | 16.2175 | 16.15 | 17.0849 | 17.1 | 16.385 | 16 | 14.8333 | 18.0967 | 15.6 | 17.0667 | 15.5333 | 16.165
| Median |  | 5.15 | 15.665 | 16.25 | 16.25 | 14.3 | 17 | 16.5456 | 16.875 | 16.975 | 15.6 | 16.575 | 15.525 | 14.715 | 16.8 | 16.2078 | 16.15 | 17.0485 | 17.1 | 16.295 | 16 | 14.85 | 18.295 | 15.45 | 17.05 | 15.55 | 16.27
| Std Dev. |  | 0.69498 | 0.51168 | 0.26583 | 0.24014 | 0.57619 | 0.41312 | 0.24836 | 0.23022 | 0.16253 | 0.42387 | 0.063246 | 0.21775 | 0.11587 | 0.31252 | 0.15544 | 0.054772 | 0.12203 | 0.33466 | 0.30885 | 0 | 0.26583 | 0.68363 | 0.54406 | 0.26583 | 0.18619 | 0.37602
| Rel.Std.Dev. |  | 13.76% | 3.257% | 1.638% | 1.484% | 4.001% | 2.435% | 1.505% | 1.362% | 0.9594% | 2.688% | 0.381% | 1.41% | 0.789% | 1.858% | 0.9585% | 0.3391% | 0.7142% | 1.957% | 1.885% | 0% | 1.792% | 3.778% | 3.488% | 1.558% | 1.199% | 2.326%
| PDM3 |  | -68.87% | -3.152% | 0.06336% | -0.2448% | -11.24% | 4.584% | 1.689% | 4.173% | 4.43% | -2.813% | 2.324% | -4.817% | -9.47% | 3.659% | -0.03415% | -0.4503% | 5.312% | 5.406% | 0.9982% | -1.375% | -8.566% | 11.55% | -3.841% | 5.2% | -4.251% | -0.3578%
| Z-Score (Absolute) |  | 11.03 | 0.5 | 0.02 | 0.03 | 1.8 | 0.74 | 0.28 | 0.67 | 0.72 | 0.45 | 0.38 | 0.77 | 1.51 | 0.59 | 0 | 0.07 | 0.86 | 0.87 | 0.17 | NA | 1.37 | 1.86 | 0.61 | 0.84 | 0.68 | 0.05
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 555b (Certified Value 0.011 wt.%) | [28 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | 1.28 | 0.01 | 0.02 | 0.01 | 0.017 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.027 | 0.007 | 0.018149 | 0.02 | 0.012862 | 0.02 | 0.01 | 0.01 | 0.0123 | 0.01 | 0.01 | 0.02 |  |  |  |  | 1
| 1 | 2 | 1.3 | 0.011 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.03 | 0.0269 | 0.007 | 0.018333 | 0.02 | 0.013604 | 0.01 | 0.01 | 0.01 | 0.0114 | 0.01 | 0.01 | 0.02 |  |  |  |  | 17
| 1 | 3 | 1.22 | 0.011 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.03 | 0.0271 | 0.007 | 0.018594 | 0.02 | 0.013379 | 0.01 | 0.01 | 0.01 | 0.0112 | 0.01 | 0.01 | 0.02 |  |  |  |  | 16
| 1 | 4 | 1.22 | 0.011 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.03 | 0.0277 | 0.007 | 0.020294 | 0.02 | 0.013673 | 0.02 | 0.01 | 0.02 | 0.0121 | 0.01 | 0.01 | 0.02 |  |  |  |  | 0.010713
| 1 | 5 | 1.22 | 0.012 | 0.02 | 0.01 | 0.019 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0266 | 0.007 | 0.020453 | 0.02 | 0.013854 | 0.01 | 0.01 | 0.01 | 0.0115 | 0.01 | 0.01 | 0.02 |  |  |  |  | 180
| 1 | 6 | 1.26 | 0.011 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0269 | 0.007 | 0.018143 | 0.02 | 0.013408 | 0.01 | 0.01 | 0.01 | 0.0107 | 0.01 | 0.01 | 0.02
| Mean |  | 1.25 | 0.011 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.025 | 0.027033 | 0.007 | 0.018995 | 0.02 | 0.013463 | 0.013333 | 0.01 | 0.011667 | 0.011533 | 0.01 | 0.01 | 0.02
| Median |  | 1.24 | 0.011 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.025 | 0.02695 | 0.007 | 0.018464 | 0.02 | 0.013506 | 0.01 | 0.01 | 0.01 | 0.01145 | 0.01 | 0.01 | 0.02
| Std Dev. |  | 0.035214 | 0.00063246 | 0 | 0 | 0.00063246 | 0 | 0 | 0 | 0 | 0 | 0.0054772 | 0.00036697 | 0 | 0.0010819 | 0 | 0.00034294 | 0.005164 | 0 | 0.0040825 | 0.00058878 | 0 | 0 | 0
| Rel.Std.Dev. |  | 2.817% | 5.75% | 0% | 0% | 3.514% | 0% | 0% | 0% | 0% | 0% | 21.91% | 1.357% | 0% | 5.696% | 0% | 2.547% | 38.73% | 0% | 34.99% | 5.105% | 0% | 0% | 0%
| PDM3 |  | 11568% | 2.675% | 86.68% | -6.659% | 68.01% | -6.659% | -6.659% | -6.659% | -6.659% | -6.659% | 133.4% | 152.3% | -34.66% | 77.3% | 86.68% | 25.67% | 24.45% | -6.659% | 8.897% | 7.653% | -6.659% | -6.659% | 86.68%
| Z-Score (Absolute) |  | 544.64 | 0.23 | 3.72 | 0.67 | 2.84 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 5.92 | 6.82 | 1.81 | 3.28 | 3.72 | 0.85 | 0.79 | 0.67 | 0.06 | 0 | 0.67 | 0.67 | 3.72
| Analytical results for Nb in OREAS 555b (Certified Value 0.094 ppm) | [28 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 28 | [8 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [8 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [8 empty] | 3
| 1 | 1 | < 5 | 0.1 | < 0.1 | 0.13 | 0.07 | < 0.05 | 0.1 | 0.1 | 0.06 | 0.11 | < 0.2 | < 0.2 | 0.15255 | 2 | < 0.1 | 0.1 | 0.13 | 0.09 | 0.14 | [8 empty] | 1
| 1 | 2 | < 5 | 0.11 | < 0.1 | 0.14 | 0.08 | < 0.05 | 0.1 | 0.1 | 0.07 | 0.12 | < 0.2 | < 0.2 | 0.15795 | 1 | < 0.1 | 0.1 | 0.1 | 0.09 | 0.12 | [8 empty] | 9
| 1 | 3 | < 5 | 0.1 | < 0.1 | 0.13 | 0.07 | < 0.05 | 0.1 | 0.09 | 0.06 | 0.11 | < 0.2 | < 0.2 | 0.1619 | 2 | < 0.1 | 0.09 | 0.09 | 0.09 | 0.13 | [8 empty] | 16
| 1 | 4 | < 5 | 0.09 | < 0.1 | 0.14 | 0.07 | 0.05 | 0.1 | 0.09 | 0.06 | 0.12 | < 0.2 | < 0.2 | 0.15309 | 1 | < 0.1 | 0.08 | 0.12 | 0.08 | 0.12 | [8 empty] | 0.093611
| 1 | 5 | < 5 | 0.09 | < 0.1 | 0.15 | 0.07 | < 0.05 | 0.09 | 0.1 | 0.05 | 0.12 | < 0.2 | < 0.2 | 0.15045 | 2 | < 0.1 | 0.09 | 0.06 | 0.1 | 0.12 | [8 empty] | 181
| 1 | 6 | < 5 | 0.1 | < 0.1 | 0.13 | 0.06 | < 0.05 | 0.09 | 0.1 | 0.05 | 0.11 | < 0.2 | < 0.2 | 0.15243 | 1 | < 0.1 | 0.08 | 0.08 | 0.09 | 0.12
| Mean |  |  | 0.098333 |  | 0.13667 | 0.07 | 0.05 | 0.096667 | 0.096667 | 0.058333 | 0.115 |  |  | 0.15473 | 1.5 |  | 0.09 | 0.096667 | 0.09 | 0.125
| Median |  |  | 0.1 |  | 0.135 | 0.07 | 0.05 | 0.1 | 0.1 | 0.06 | 0.115 |  |  | 0.15282 | 1.5 |  | 0.09 | 0.095 | 0.09 | 0.12
| Std Dev. |  |  | 0.0075277 |  | 0.008165 | 0.0063246 |  | 0.005164 | 0.005164 | 0.0075277 | 0.0054772 |  |  | 0.0043107 | 0.54772 |  | 0.0089443 | 0.02582 | 0.0063246 | 0.0083666
| Rel.Std.Dev. |  |  | 7.655% |  | 5.974% | 9.035% |  | 5.342% | 5.342% | 12.9% | 4.763% |  |  | 2.786% | 36.51% |  | 9.938% | 26.71% | 7.027% | 6.693%
| PDM3 |  |  | 5.045% |  | 45.99% | -25.22% | -46.59% | 3.264% | 3.264% | -37.69% | 22.85% |  |  | 65.29% | 1502% |  | -3.858% | 3.264% | -3.858% | 33.53%
| Z-Score (Absolute) |  | 72.03 | 0.05 | 1.4 | 1.2 | 0.8 | 2.02 | 0 | 0 | 1.15 | 0.55 | 0.1 | 0.1 | 1.74 | NA | 1.4 | 0.2 | 0 | 0.2 | 0.85
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 555b (Certified Value 23 ppm) | [28 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 1
| 1 | 1 | 24.9 | 27.1 | 18.4 | 11.3305 | 20.5028 | 25.95 | [21 empty] | 1
| 1 | 2 | 24.36 | 25.3 | 18.5 | 11.202 | 20.6074 | 27.04 | [21 empty] | 10
| 1 | 3 | 23.321 | 24.7 | 19.5 | 11.4268 | 20.7633 | 25.65 | [21 empty] | 16
| 1 | 4 | 23.756 | 24.7 | 18.6 | 11.6423 | 19.2562 | 26.44 | [21 empty] | 22.9844
| 1 | 5 | 24.041 | 25.8 | 18.6 | 10.8675 | 20.3888 | 26.23 | [21 empty] | 182
| 1 | 6 | 24.226 | 24.7 | 18.8 | 10.6111 | 20.3795 | 27.02
| Mean |  | 24.1007 | 25.3833 | 18.7333 | 11.18 | 20.3163 | 26.3883
| Median |  | 24.1335 | 25 | 18.6 | 11.2662 | 20.4458 | 26.335
| Std Dev. |  | 0.53891 | 0.95167 | 0.39833 | 0.37924 | 0.53894 | 0.56368
| Rel.Std.Dev. |  | 2.236% | 3.749% | 2.126% | 3.392% | 2.653% | 2.136%
| PDM3 |  | 4.857% | 10.44% | -18.5% | -51.36% | -11.61% | 14.81%
| Z-Score (Absolute) |  | 0.38 | 0.64 | 0.7 | 2.24 | 0.38 | 0.85
| Analytical results for Ni in OREAS 555b (Certified Value 214 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 0
| 1 | 1 | 186.1 | 200.6 | 215 | 233.3 | 228 | 217 | 213.4446 | 203 | 213 | 208 | 206 | 211 | 205 | 199 | 214.1217 | 193.08 | 227 | 220.5686 | 222 | 207 | 200.4 | 131 | 217 | 224.9 | 219.3 | 222.5 |  | 1
| 1 | 2 | 184.1 | 229 | 207 | 234.4 | 224 | 223 | 216.612 | 206 | 231 | 207 | 209 | 213 | 200 | 202 | 215.8283 | 187.13 | 234 | 219.9371 | 231 | 205 | 206.2 | 142 | 206 | 228.2 | 213 | 221.8 |  | 40
| 1 | 3 | 186 | 225.4 | 215 | 233 | 212 | 225 | 213.7037 | 211 | 225 | 208 | 211 | 213 | 204 | 218 | 216.5432 | 190.96 | 232 | 221.096 | 237 | 202 | 206.1 | 140 | 205 | 223.5 | 211.4 | 219.1 |  | 16
| 1 | 4 | 180 | 222.8 | 210 | 243.1 | 212 | 231 | 211.4293 | 216 | 224 | 208 | 210 | 220 | 208 | 196 | 214.126 | 190.57 | 224 | 220.8664 | 233 | 200 | 205.5 | 141 | 202 | 218.3 | 216.8 | 221.8 |  | 213.8519
| 1 | 5 | 186.6 | 243.4 | 216 | 217.5 | 204 | 229 | 214.0126 | 211 | 220 | 220 | 215 | 215 | 201 | 194 | 217.4283 | 190.95 | 225 | 222.3735 | 222 | 209 | 204.4 | 147 | 204 | 224.5 | 216.9 | 222.9 |  | 183
| 1 | 6 | 186.8 | 231.2 | 220 | 238.5 | 212 | 237 | 211.3471 | 214 | 216 | 213 | 209 | 218 | 206 | 202 | 215.669 | 191.69 | 224 | 227.511 | 230 | 212 | 197.1 | 152 | 202 | 234.2 | 216.5 | 220.6
| Mean |  | 184.9333 | 225.4 | 213.8333 | 233.3 | 215.3333 | 227 | 213.4249 | 210.1667 | 221.5 | 210.6667 | 210 | 215 | 204 | 201.8333 | 215.6194 | 190.73 | 227.6667 | 222.0588 | 229.1667 | 205.8333 | 203.2833 | 142.1667 | 206 | 225.6 | 215.65 | 221.45
| Median |  | 186.05 | 227.2 | 215 | 233.85 | 212 | 227 | 213.5742 | 211 | 222 | 208 | 209.5 | 214 | 204.5 | 200.5 | 215.7486 | 190.955 | 226 | 220.9812 | 230.5 | 206 | 204.95 | 141.5 | 204.5 | 224.7 | 216.65 | 221.8
| Std Dev. |  | 2.5997 | 14.0883 | 4.6224 | 8.6489 | 8.9144 | 6.9282 | 1.9432 | 4.8751 | 6.5345 | 5.0465 | 2.9665 | 3.4059 | 3.0332 | 8.5421 | 1.3149 | 1.9764 | 4.3205 | 2.7891 | 6.047 | 4.446 | 3.7188 | 7.0828 | 5.6214 | 5.2938 | 2.8988 | 1.3925
| Rel.Std.Dev. |  | 1.406% | 6.25% | 2.162% | 3.707% | 4.14% | 3.052% | 0.9105% | 2.32% | 2.95% | 2.395% | 1.413% | 1.584% | 1.487% | 4.232% | 0.6098% | 1.036% | 1.898% | 1.256% | 2.639% | 2.16% | 1.829% | 4.982% | 2.729% | 2.347% | 1.344% | 0.6288%
| PDM3 |  | -13.52% | 5.4% | -0.008661% | 9.094% | 0.6928% | 6.148% | -0.1997% | -1.723% | 3.576% | -1.489% | -1.801% | 0.5369% | -4.607% | -5.62% | 0.8265% | -10.81% | 6.46% | 3.838% | 7.161% | -3.75% | -4.942% | -33.52% | -3.672% | 5.494% | 0.8408% | 3.553%
| Z-Score (Absolute) |  | 2.34 | 0.87 | 0.05 | 1.5 | 0.07 | 1 | 0.08 | 0.34 | 0.56 | 0.3 | 0.35 | 0.05 | 0.83 | 1 | 0.1 | 1.88 | 1.05 | 0.61 | 1.17 | 0.68 | 0.88 | 5.73 | 0.67 | 0.89 | 0.1 | 0.56
| Analytical results for P in OREAS 555b (Certified Value 0.02 wt.%) | [28 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02 | 0.0177 | 0.0501 | 0.0201 | 0.036 | 0.037 | 0.019 | 0.02 | 0.019 | 0.019 | 0.02 | 0.0561 | 0.1536 | 0.020529 | 0.0248 | 0.020183 | 0.02 | 0.02 | 0.018 | 0.02 | 0.0232 | 0.0218 | 0.019 | 0.0199 |  |  |  | 1
| 1 | 2 | 0.02 | 0.0196 | 0.0502 | 0.0202 | 0.038 | 0.035 | 0.019 | 0.021 | 0.019 | 0.019 | 0.019 | 0.0576 | 0.1689 | 0.018904 | 0.0249 | 0.020976 | 0.03 | 0.02 | 0.018 | 0.022 | 0.0226 | 0.0214 | 0.02 | 0.0196 |  |  |  | #N/A
| 1 | 3 | 0.02 | 0.0195 | 0.0516 | 0.0208 | 0.037 | 0.033 | 0.02 | 0.02 | 0.019 | 0.019 | 0.02 | 0.0575 | 0.1457 | 0.020009 | 0.0244 | 0.020269 | 0.02 | 0.02 | 0.018 | 0.02 | 0.0215 | 0.0214 | 0.019 | 0.0195 |  |  |  | 16
| 1 | 4 | 0.02 | 0.02 | 0.05 | 0.0205 | 0.036 | 0.034 | 0.02 | 0.021 | 0.02 | 0.019 | 0.02 | 0.0568 | 0.1383 | 0.019376 | 0.0244 | 0.019965 | 0.03 | 0.02 | 0.018 | 0.021 | 0.0236 | 0.0215 | 0.02 | 0.0187 |  |  |  | 0.020017
| 1 | 5 | 0.02 | 0.021 | 0.0522 | 0.0194 | 0.036 | 0.033 | 0.02 | 0.021 | 0.02 | 0.019 | 0.019 | 0.0566 | 0.1488 | 0.01988 | 0.0253 | 0.020919 | 0.02 | 0.02 | 0.017 | 0.021 | 0.0234 | 0.0217 | 0.019 | 0.0196 |  |  |  | 184
| 1 | 6 | 0.02 | 0.021 | 0.0519 | 0.0198 | 0.036 | 0.038 | 0.02 | 0.02 | 0.019 | 0.019 | 0.019 | 0.0574 | 0.1436 | 0.021055 | 0.0253 | 0.020552 | 0.02 | 0.02 | 0.018 | 0.021 | 0.0214 | 0.0218 | 0.019 | 0.02
| Mean |  | 0.02 | 0.0198 | 0.051 | 0.020133 | 0.0365 | 0.035 | 0.019667 | 0.0205 | 0.019333 | 0.019 | 0.0195 | 0.057 | 0.14982 | 0.019959 | 0.02485 | 0.020477 | 0.023333 | 0.02 | 0.017833 | 0.020833 | 0.022617 | 0.0216 | 0.019333 | 0.01955
| Median |  | 0.02 | 0.0198 | 0.0509 | 0.02015 | 0.036 | 0.0345 | 0.02 | 0.0205 | 0.019 | 0.019 | 0.0195 | 0.0571 | 0.14725 | 0.019945 | 0.02485 | 0.020411 | 0.02 | 0.02 | 0.018 | 0.021 | 0.0229 | 0.0216 | 0.019 | 0.0196
| Std Dev. |  | 0 | 0.0012215 | 0.001006 | 0.00049666 | 0.00083666 | 0.0020976 | 0.0005164 | 0.00054772 | 0.0005164 | 0 | 0.00054772 | 0.00059666 | 0.010656 | 0.00077294 | 0.00040373 | 0.00041048 | 0.005164 | 0 | 0.00040825 | 0.00075277 | 0.00096419 | 0.00018974 | 0.0005164 | 0.00045935
| Rel.Std.Dev. |  | 0% | 6.169% | 1.973% | 2.467% | 2.292% | 5.993% | 2.626% | 2.672% | 2.671% | 0% | 2.809% | 1.047% | 7.113% | 3.873% | 1.625% | 2.005% | 22.13% | 0% | 2.289% | 3.613% | 4.263% | 0.8784% | 2.671% | 2.35%
| PDM3 |  | -0.08506% | -1.084% | 154.8% | 0.581% | 82.34% | 74.85% | -1.75% | 2.413% | -3.416% | -5.081% | -2.583% | 184.8% | 648.4% | -0.2899% | 24.14% | 2.3% | 16.57% | -0.08506% | -10.91% | 4.078% | 12.99% | 7.908% | -3.416% | -2.333%
| Z-Score (Absolute) |  | 0.21 | 0.35 | 21.29 | 0.12 | 11.23 | 10.19 | 0.44 | 0.13 | 0.67 | 0.91 | 0.56 | 25.45 | 89.84 | 0.24 | 3.15 | 0.12 | 2.1 | 0.21 | 1.71 | 0.37 | 1.6 | 0.9 | 0.67 | 0.52
| Analytical results for Pb in OREAS 555b (Certified Value 10.9 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 2
| 1 | 1 | 28.2 | 8.51 | 12.1 | 11 | 11.2 | 10.6 | 10.1865 | 9.5 | 11.1 | 9.9 | 10.4 | 9.8 | 13.72 | 10.9 | 10.9016 | 15.39 | 12 | 11.22 | 10 | 11.5 | 15 | 9 | 13.4 | 11.7 | 10.8 | 11.5 |  | 1
| 1 | 2 | 27.3 | 8.43 | 11 | 10.9 | 11 | 11.2 | 9.9494 | 9.8 | 10.8 | 9.8 | 10.4 | 9.9 | 13.31 | 11.5 | 10.5402 | 15.27 | 12.2 | 12.1663 | 11 | 11.4 | 15 | 9.4 | 13.2 | 11.5 | 10.9 | 11 |  | 11
| 1 | 3 | 29.3 | 9.08 | 11.2 | 10.7 | 10.5 | 14.3 | 10.0001 | 9.7 | 11 | 10 | 10 | 9.7 | 13.45 | 11.6 | 11.2563 | 15.92 | 12 | 11.5466 | 10 | 11.6 | 15 | 9.3 | 12.7 | 11.3 | 10.6 | 11.1 |  | 16
| 1 | 4 | 28.2 | 8.93 | 11.9 | 10.8 | 10.3 | 11.4 | 10.2988 | 9.7 | 11 | 9.8 | 10.4 | 9.8 | 13.44 | 11 | 10.7474 | 15.18 | 12 | 11.3253 | 10 | 11.5 | 16 | 9.5 | 12.7 | 11.3 | 11 | 10.9 |  | 10.9141
| 1 | 5 | 28.2 | 9.98 | 11.8 | 10.7 | 10.9 | 11 | 10.5533 | 9.9 | 11 | 10.2 | 10.2 | 9.5 | 13.62 | 11 | 10.6548 | 15.8 | 12 | 11.9189 | 16 | 11.5 | 15 | 9.3 | 11.3 | 11.6 | 10.8 | 11.1 |  | 185
| 1 | 6 | 28.2 | 8.91 | 11.8 | 10.9 | 10.6 | 11.2 | 10.2743 | 9.8 | 11.2 | 10 | 10.3 | 9.4 | 13.67 | 11.1 | 11.0972 | 15.24 | 11.9 | 11.4886 | 14 | 11.4 | 14 | 9.4 | 11.9 | 11.6 | 10.9 | 10.9
| Mean |  | 28.2333 | 8.9733 | 11.6333 | 10.8333 | 10.75 | 11.6167 | 10.2104 | 9.7333 | 11.0167 | 9.95 | 10.2833 | 9.6833 | 13.535 | 11.1833 | 10.8662 | 15.4667 | 12.0167 | 11.6109 | 11.8333 | 11.4833 | 15 | 9.3167 | 12.5333 | 11.5 | 10.8333 | 11.0833
| Median |  | 28.2 | 8.92 | 11.8 | 10.85 | 10.75 | 11.2 | 10.2304 | 9.75 | 11 | 9.95 | 10.35 | 9.75 | 13.535 | 11.05 | 10.8245 | 15.33 | 12 | 11.5176 | 10.5 | 11.5 | 15 | 9.35 | 12.7 | 11.55 | 10.85 | 11.05
| Std Dev. |  | 0.63456 | 0.55507 | 0.43205 | 0.12111 | 0.33912 | 1.3423 | 0.22016 | 0.13663 | 0.13292 | 0.15166 | 0.16021 | 0.19408 | 0.15909 | 0.29269 | 0.27273 | 0.31456 | 0.098319 | 0.36243 | 2.5626 | 0.075277 | 0.63246 | 0.17224 | 0.79666 | 0.16733 | 0.13663 | 0.22286
| Rel.Std.Dev. |  | 2.248% | 6.186% | 3.714% | 1.118% | 3.155% | 11.55% | 2.156% | 1.404% | 1.206% | 1.524% | 1.558% | 2.004% | 1.175% | 2.617% | 2.51% | 2.034% | 0.8182% | 3.121% | 21.66% | 0.6555% | 4.216% | 1.849% | 6.356% | 1.455% | 1.261% | 2.011%
| PDM3 |  | 158.7% | -17.78% | 6.59% | -0.7396% | -1.503% | 6.438% | -6.447% | -10.82% | 0.9402% | -8.833% | -5.779% | -11.28% | 24.01% | 2.467% | -0.4381% | 41.71% | 10.1% | 6.385% | 8.423% | 5.216% | 37.44% | -14.64% | 14.84% | 5.369% | -0.7396% | 1.551%
| Z-Score (Absolute) |  | 14.46 | 1.78 | 0.46 | 0.21 | 0.28 | 0.45 | 0.74 | 1.14 | 0.06 | 0.96 | 0.67 | 1.18 | 2.07 | 0.08 | 0.18 | 3.69 | 0.79 | 0.44 | NA | 0.34 | 3.3 | 1.49 | 1.22 | 0.35 | 0.21 | 0
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 555b (Indicative Value < 10 ppb) | [28 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 04 | 06 | 07 | 15 | 26 | [22 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [22 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 01g | [22 empty] | 1
| 1 | 1 | < 10 | 18 | < 20 | < 10 | < 10 | [22 empty] | 1
| 1 | 2 | < 10 | 15 | < 20 | < 10 | < 10 | [22 empty] | 26
| 1 | 3 | < 10 | 16 | < 20 | < 10 | < 10 | [22 empty] | 16
| 1 | 4 | < 10 | 17 | < 20 | < 10 | < 10 | [22 empty] | < 10
| 1 | 5 | < 10 | 19 | < 20 | < 10 | < 10 | [22 empty] | 32
| 1 | 6 | < 10 | 16 | < 20 | < 10 | < 10
| Mean |  |  | 16.8333
| Median |  |  | 16.5
| Std Dev. |  |  | 1.472
| Rel.Std.Dev. |  |  | 8.744%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Pr in OREAS 555b (Certified Value 5.43 ppm) | [28 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 2
| 1 | 1 | 6.314 | 6.7 | 4.59 | 2.9414 | 5.1625 | 7.15 | [21 empty] | 1
| 1 | 2 | 6.307 | 6.4 | 4.72 | 2.9195 | 5.223 | 7.38 | [21 empty] | 12
| 1 | 3 | 6.009 | 6.2 | 4.87 | 2.9089 | 5.3214 | 7.05 | [21 empty] | 16
| 1 | 4 | 6.212 | 6 | 4.73 | 3.0181 | 4.9422 | 7.25 | [21 empty] | 5.4327
| 1 | 5 | 6.293 | 6.4 | 4.66 | 2.8638 | 5.0947 | 7.17 | [21 empty] | 186
| 1 | 6 | 6.313 | 6.1 | 4.76 | 2.7593 | 5.2346 | 7.33
| Mean |  | 6.2413 | 6.3 | 4.7217 | 2.9018 | 5.1631 | 7.2217
| Median |  | 6.3 | 6.3 | 4.725 | 2.9142 | 5.1927 | 7.21
| Std Dev. |  | 0.1202 | 0.25298 | 0.094534 | 0.086251 | 0.13208 | 0.12238
| Rel.Std.Dev. |  | 1.926% | 4.016% | 2.002% | 2.972% | 2.558% | 1.695%
| PDM3 |  | 14.89% | 15.97% | -13.09% | -46.59% | -4.963% | 32.93%
| Z-Score (Absolute) |  | 0.46 | 0.51 | 0.84 | 2.39 | 0.46 | 1.3
| Analytical results for Pt in OREAS 555b (Indicative Value < 5 ppb) | [28 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 06 | 07 | 15 | 25 | 26 | [21 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.1g | 01g | [21 empty] | 2
| 1 | 1 | < 5 | < 10 | 9.82 | < 5 | < 1 | < 5 | [21 empty] | 1
| 1 | 2 | < 5 | < 10 | < 5 | < 5 | < 1 | 5 | [21 empty] | 22
| 1 | 3 | < 5 | < 10 | < 5 | < 5 | < 1 | < 5 | [21 empty] | 16
| 1 | 4 | < 5 | < 10 | < 5 | < 5 | 1 | < 5 | [21 empty] | < 5
| 1 | 5 | < 5 | < 10 | < 5 | < 5 | 3 | 7 | [21 empty] | 30
| 1 | 6 | < 5 | < 10 | 14.54 | < 5 | 1 | 6
| Mean |  |  |  | 12.18 |  | 1.6667 | 6
| Median |  |  |  | 12.18 |  | 1 | 6
| Std Dev. |  |  |  | 3.3375 |  | 1.1547 | 1
| Rel.Std.Dev. |  |  |  | 27.4% |  | 69.28% | 16.67%
PDM3
| Z-Score (Absolute) |  | 0.47 | 0.88 | 1.27 | 0.47 | 1.24 | 0.47
| Analytical results for Rb in OREAS 555b (Certified Value 6.5 ppm) | [28 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 22 | 23 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [10 empty] | 2
| 1 | 1 | 8.28 | 8.9 | 5.3 | 5.1 | 7 | 5.8 | 5.8 | 5.9 | 4.59 | 7.9128 | 11.8 | 6.7604 | 6.2 | 6.69 | 13.8 | 17.55 | 7.2 | [10 empty] | 1
| 1 | 2 | 8.13 | 8.8 | 5.6 | 5.3 | 7.6 | 5.5 | 5.9 | 6.1 | 4.7 | 7.1977 | 11.6 | 6.5799 | 6.1 | 6.83 | 17.1 | 16.86 | 7.1 | [10 empty] | 13
| 1 | 3 | 8.21 | 8.7 | 5.4 | 5.3 | 7.5 | 5.5 | 5.9 | 5.9 | 4.85 | 7.738 | 12.7 | 6.5324 | 6.1 | 6.78 | 16.7 | 16.14 | 6.9 | [10 empty] | 16
| 1 | 4 | 7.79 | 8.2 | 5.5 | 5.4 | 7.2 | 5.7 | 6 | 6 | 4.46 | 7.2859 | 11.5 | 6.3798 | 6.2 | 6.87 | 15.7 | 16.76 | 6.8 | [10 empty] | 6.4953
| 1 | 5 | 7.75 | 8.6 | 5.7 | 5.6 | 7.6 | 5.7 | 5.9 | 5.7 | 4.58 | 7.7589 | 11.8 | 6.2774 | 5.9 | 7.02 | 11.3 | 16.86 | 7 | [10 empty] | 187
| 1 | 6 | 8.34 | 8.2 | 5.7 | 5.3 | 7.3 | 5.3 | 5.8 | 5.7 | 4.66 | 7.9361 | 11.4 | 6.2674 | 5.9 | 7.05 | 10.5 | 17.12 | 7.1
| Mean |  | 8.0833 | 8.5667 | 5.5333 | 5.3333 | 7.3667 | 5.5833 | 5.8833 | 5.8833 | 4.64 | 7.6382 | 11.8 | 6.4662 | 6.0667 | 6.8733 | 14.1833 | 16.8817 | 7.0167
| Median |  | 8.17 | 8.65 | 5.55 | 5.3 | 7.4 | 5.6 | 5.9 | 5.9 | 4.625 | 7.7485 | 11.7 | 6.4561 | 6.1 | 6.85 | 14.75 | 16.86 | 7.05
| Std Dev. |  | 0.25296 | 0.30111 | 0.1633 | 0.1633 | 0.24221 | 0.18348 | 0.075277 | 0.16021 | 0.13161 | 0.31837 | 0.46904 | 0.19313 | 0.13663 | 0.13924 | 2.7989 | 0.46245 | 0.1472
| Rel.Std.Dev. |  | 3.129% | 3.515% | 2.951% | 3.062% | 3.288% | 3.286% | 1.28% | 2.723% | 2.836% | 4.168% | 3.975% | 2.987% | 2.252% | 2.026% | 19.73% | 2.739% | 2.098%
| PDM3 |  | 24.45% | 31.89% | -14.81% | -17.89% | 13.42% | -14.04% | -9.422% | -9.422% | -28.56% | 17.6% | 81.67% | -0.4481% | -6.599% | 5.82% | 118.4% | 159.9% | 8.027%
| Z-Score (Absolute) |  | 0.67 | 0.94 | 0.75 | 0.86 | 0.27 | 0.72 | 0.55 | 0.55 | 1.24 | 0.43 | 2.75 | 0.23 | 0.45 | 0 | 4.07 | 5.58 | 0.08
| Analytical results for Re in OREAS 555b (Certified Value 0.027 ppm) | [28 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 20 | 22 | 25 | 26 | 28 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [11 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.1g | 01g | 0.5g | [11 empty] | 3
| 1 | 1 | 0.027 | 0.025 | 0.03 | 0.023 | 0.028 | 0.027 | 0.025 | 0.026 | < 0.05 | < 0.05 | 0.025873 | 0.027 | 0.025 | 0.026 | 0.031 | 0.025 | [11 empty] | 1
| 1 | 2 | 0.025 | 0.025 | 0.031 | 0.023 | 0.028 | 0.025 | 0.028 | 0.026 | < 0.05 | < 0.05 | 0.02487 | 0.03 | 0.026 | 0.022 | 0.031 | 0.026 | [11 empty] | 14
| 1 | 3 | 0.027 | 0.025 | 0.03 | 0.026 | 0.03 | 0.026 | 0.024 | 0.027 | < 0.05 | < 0.05 | 0.024223 | 0.03 | 0.026 | 0.024 | 0.031 | 0.024 | [11 empty] | 16
| 1 | 4 | 0.031 | 0.024 | 0.027 | 0.027 | 0.029 | 0.024 | 0.025 | 0.026 | < 0.05 | < 0.05 | 0.023277 | 0.031 | 0.027 | 0.022 | 0.029 | 0.025 | [11 empty] | 0.026561
| 1 | 5 | 0.025 | 0.024 | 0.031 | 0.024 | 0.03 | 0.028 | 0.027 | 0.027 | < 0.05 | < 0.05 | 0.026067 | 0.027 | 0.026 | 0.022 | 0.03 | 0.023 | [11 empty] | 188
| 1 | 6 | 0.033 | 0.025 | 0.034 | 0.026 | 0.028 | 0.025 | 0.029 | 0.026 | < 0.05 | < 0.05 | 0.024421 | 0.024 | 0.027 | 0.022 | 0.031 | 0.024
| Mean |  | 0.028 | 0.024667 | 0.0305 | 0.024833 | 0.028833 | 0.025833 | 0.026333 | 0.026333 |  |  | 0.024789 | 0.028167 | 0.026167 | 0.023 | 0.0305 | 0.0245
| Median |  | 0.027 | 0.025 | 0.0305 | 0.025 | 0.0285 | 0.0255 | 0.026 | 0.026 |  |  | 0.024645 | 0.0285 | 0.026 | 0.022 | 0.031 | 0.0245
| Std Dev. |  | 0.0032863 | 0.0005164 | 0.0022583 | 0.0017224 | 0.00098319 | 0.001472 | 0.0019664 | 0.0005164 |  |  | 0.0010543 | 0.0026394 | 0.00075277 | 0.0016733 | 0.00083666 | 0.0010488
| Rel.Std.Dev. |  | 11.74% | 2.094% | 7.404% | 6.936% | 3.41% | 5.698% | 7.467% | 1.961% |  |  | 4.253% | 9.371% | 2.877% | 7.275% | 2.743% | 4.281%
| PDM3 |  | 5.417% | -7.132% | 14.83% | -6.505% | 8.555% | -2.74% | -0.8575% | -0.8575% |  |  | -6.674% | 6.045% | -1.485% | -13.41% | 14.83% | -7.76%
| Z-Score (Absolute) |  | 1.06 | 0.71 | 2.38 | 0.62 | 1.5 | 0.09 | 0.18 | 0.18 | 0.53 | 0.53 | 0.64 | 1.15 | 0.09 | 1.59 | 2.38 | 0.79
| Analytical results for S in OREAS 555b (Certified Value 4.31 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 3
| 1 | 1 | 3.54 | 3.94 | 4.07 | 4.4216 | 4.172 | 4.51 | 4.18 | 4.38 | 4.35 | 4.22 | 4.31 | 4.32 | 5.04 | 4.3752 | 3.9103 | 4.04 | 4.3773 | 4.2159 | 4.34 | 4.47 | 4.15 | 4.35 | 4.3 | 4.4452 | 4.33 | 4.69 |  | 1
| 1 | 2 | 3.39 | 4.4 | 4 | 4.4307 | 4.195 | 4.5 | 4.24 | 4.48 | 4.38 | 4.2 | 4.32 | 4.35 | 4.88 | 4.1416 | 3.9014 | 4.2 | 4.5131 | 4.2284 | 4.4 | 4.48 | 4.22 | 4.58 | 4.18 | 4.3747 | 4.27 | 4.6 |  | 28
| 1 | 3 | 3.59 | 4.36 | 4 | 4.5775 | 4.097 | 4.39 | 4.27 | 4.4 | 4.35 | 4.2 | 4.33 | 4.34 | 5.25 | 4.2615 | 3.9568 | 4.16 | 4.4625 | 4.285 | 4.46 | 4.37 | 4.23 | 4.28 | 4.06 | 4.476 | 4.31 | 4.56 |  | 16
| 1 | 4 | 3.59 | 4.35 | 3.98 | 4.565 | 3.995 | 4.49 | 4.33 | 4.53 | 4.4 | 4.23 | 4.31 | 4.32 | 4.76 | 4.3226 | 3.9981 | 4.17 | 4.4415 | 4.2592 | 4.41 | 4.41 | 4.24 | 4.32 | 4.26 | 4.4629 | 4.33 | 4.62 |  | 4.3112
| 1 | 5 | 3.49 | 4.57 | 4.1 | 4.3818 | 4.019 | 4.41 | 4.37 | 4.5 | 4.38 | 4.19 | 4.24 | 4.36 | 4.88 | 4.23 | 3.9917 | 4.08 | 4.3844 | 4.2641 | 4.44 | 4.49 | 4.19 | 4.41 | 4.11 | 4.4635 | 4.29 | 4.61 |  | 189
| 1 | 6 | 3.43 | 4.6 | 4.06 | 4.3502 | 4.122 | 4.62 | 4.27 | 4.31 | 4.44 | 4.25 | 4.2 | 4.37 | 4.84 | 4.2139 | 3.9637 | 4.08 | 4.4282 | 4.3209 | 4.33 | 4.46 | 4.32 | 4.49 | 4.28 | 4.5634 | 4.32 | 4.6
| Mean |  | 3.505 | 4.37 | 4.035 | 4.4545 | 4.1 | 4.4867 | 4.2767 | 4.4333 | 4.3833 | 4.215 | 4.285 | 4.3433 | 4.9417 | 4.2575 | 3.9537 | 4.1217 | 4.4345 | 4.2623 | 4.3967 | 4.4467 | 4.225 | 4.405 | 4.1983 | 4.4643 | 4.3083 | 4.6133
| Median |  | 3.515 | 4.38 | 4.03 | 4.4261 | 4.1095 | 4.495 | 4.27 | 4.44 | 4.38 | 4.21 | 4.31 | 4.345 | 4.88 | 4.2458 | 3.9602 | 4.12 | 4.4348 | 4.2616 | 4.405 | 4.465 | 4.225 | 4.38 | 4.22 | 4.4632 | 4.315 | 4.605
| Std Dev. |  | 0.083367 | 0.2363 | 0.048062 | 0.095028 | 0.080359 | 0.082138 | 0.066833 | 0.083825 | 0.033862 | 0.022583 | 0.05244 | 0.020656 | 0.17646 | 0.082709 | 0.040339 | 0.063377 | 0.050659 | 0.038116 | 0.052409 | 0.046762 | 0.056833 | 0.11292 | 0.098065 | 0.060586 | 0.024014 | 0.04274
| Rel.Std.Dev. |  | 2.379% | 5.407% | 1.191% | 2.133% | 1.96% | 1.831% | 1.563% | 1.891% | 0.7725% | 0.5358% | 1.224% | 0.4756% | 3.571% | 1.943% | 1.02% | 1.538% | 1.142% | 0.8943% | 1.192% | 1.052% | 1.345% | 2.563% | 2.336% | 1.357% | 0.5574% | 0.9264%
| PDM3 |  | -18.7% | 1.363% | -6.408% | 3.322% | -4.9% | 4.069% | -0.8021% | 2.832% | 1.672% | -2.232% | -0.6088% | 0.7442% | 14.62% | -1.247% | -8.294% | -4.397% | 2.858% | -1.137% | 1.981% | 3.141% | -2.001% | 2.175% | -2.619% | 3.55% | -0.0676% | 7.007%
| Z-Score (Absolute) |  | 5.04 | 0.27 | 1.79 | 0.79 | 1.39 | 0.99 | 0.3 | 0.66 | 0.35 | 0.68 | 0.25 | 0.11 | 3.78 | 0.42 | 2.29 | 1.25 | 0.67 | 0.39 | 0.44 | 0.74 | 0.62 | 0.49 | 0.78 | 0.85 | 0.11 | 1.77
| Analytical results for Sb in OREAS 555b (Certified Value 2.17 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 2
| 1 | 1 | < 5 | 2.46 | < 5 | 3.03 | 2.51 | 2.09 | 2.3949 | 1.82 | 1.76 | 2.17 | 2.32 | 1.6 | < 5 | 0.83 | 0.68998 | 2.23 | 2.3596 | NA | 6 | 2 | 1.21 | 1.62 | 2.35 | 2.5 | 2.52 |  |  | 1
| 1 | 2 | < 5 | 2.45 | < 5 | 3.03 | 2.51 | 2.39 | 2.2831 | 1.88 | 1.88 | 2.14 | 2.27 | 1.56 | < 5 | 0.77 | 0.74541 | 2.19 | 2.424 | NA | 5 | 2.06 | 1.32 | 1.72 | 2.3 | 2.3 | 2.59 |  |  | 15
| 1 | 3 | < 5 | 2.66 | < 5 | 2.96 | 2.33 | 2.08 | 2.2377 | 1.91 | 1.85 | 2.11 | 2.31 | 1.49 | < 5 | 0.76 | 0.69722 | 2.12 | 2.2777 | NA | 5 | 2.04 | 1.39 | 1.64 | 2.32 | 2.3 | 2.46 |  |  | 16
| 1 | 4 | < 5 | 2.44 | < 5 | 2.99 | 2.35 | 2.2 | 2.2575 | 1.9 | 1.75 | 2.24 | 2.3 | 1.46 | < 5 | 0.74 | 0.736 | 2.12 | 2.2319 | NA | < 5 | 2.13 | 1.26 | 1.6 | 2.26 | 2.4 | 2.42 |  |  | 2.1662
| 1 | 5 | < 5 | 2.88 | < 5 | 3.08 | 2.58 | 2.24 | 2.2703 | 1.96 | 1.78 | 2.19 | 2.19 | 1.44 | < 5 | 0.8 | 0.71015 | 2.23 | 2.1998 | NA | 5 | 2 | 1.41 | 2.12 | 2.43 | 2.4 | 2.58 |  |  | 190
| 1 | 6 | < 5 | 2.65 | < 5 | 3 | 2.4 | 2.17 | 2.2706 | 1.87 | 1.96 | 2.12 | 2.09 | 1.41 | < 5 | 0.8 | 0.69757 | 2.09 | 2.2146 | NA | 5 | 2.16 | 1.36 | 1.92 | 2.42 | 2.4 | 2.37
| Mean |  |  | 2.59 |  | 3.015 | 2.4467 | 2.195 | 2.2857 | 1.89 | 1.83 | 2.1617 | 2.2467 | 1.4933 |  | 0.78333 | 0.71272 | 2.1633 | 2.2846 |  | 5.2 | 2.065 | 1.325 | 1.77 | 2.3467 | 2.3833 | 2.49
| Median |  |  | 2.555 |  | 3.015 | 2.455 | 2.185 | 2.2705 | 1.89 | 1.815 | 2.155 | 2.285 | 1.475 |  | 0.785 | 0.70386 | 2.155 | 2.2548 |  | 5 | 2.05 | 1.34 | 1.68 | 2.335 | 2.4 | 2.49
| Std Dev. |  |  | 0.17413 |  | 0.041352 | 0.10093 | 0.11397 | 0.055675 | 0.046476 | 0.081976 | 0.048751 | 0.090037 | 0.073121 |  | 0.03266 | 0.022823 | 0.06121 | 0.089433 |  | 0.44721 | 0.066858 | 0.077653 | 0.20775 | 0.067429 | 0.075277 | 0.088544
| Rel.Std.Dev. |  |  | 6.723% |  | 1.372% | 4.125% | 5.192% | 2.436% | 2.459% | 4.48% | 2.255% | 4.008% | 4.896% |  | 4.169% | 3.202% | 2.829% | 3.915% |  | 8.6% | 3.238% | 5.861% | 11.74% | 2.873% | 3.158% | 3.556%
| PDM3 |  |  | 19.56% |  | 39.18% | 12.95% | 1.33% | 5.516% | -12.75% | -15.52% | -0.2089% | 3.715% | -31.06% |  | -63.84% | -67.1% | -0.1319% | 5.467% |  | 140.1% | -4.671% | -38.83% | -18.29% | 8.331% | 10.02% | 14.95%
| Z-Score (Absolute) |  | 0.67 | 0.91 | 0.67 | 2.05 | 0.53 | 0.14 | 0.1 | 0.95 | 1.11 | 0.23 | 0 | 2.01 | 0.67 | 3.9 | 4.08 | 0.22 | 0.1 | NA | NA | 0.48 | 2.45 | 1.27 | 0.27 | 0.36 | 0.65
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 555b (Certified Value 3.3 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 2
| 1 | 1 | 4 | 3 | < 5 | 3.3 | 3.2 | 2.9 | 3.4 | 3.7 | 3.4 | 3.2 | 3.5 | 3 | 2.6213 | 4 | 3.7527 | 4.3 | 2.7 | 3.5 | 3.3 | 4 | 3.8 | 3.1 | 3.6 |  |  |  |  | 1
| 1 | 2 | 3.5 | 3.3 | < 5 | 3.3 | 3 | 3 | 3.3 | 3.7 | 3.2 | 3.3 | 3.4 | 2.9 | 2.5095 | 4 | 3.8687 | 4.6 | 2.7 | 3.2 | 3.6 | 5 | 3.9 | 2.9 | 3.5 |  |  |  |  | 1
| 1 | 3 | 3.6 | 3.4 | < 5 | 3.4 | 2.9 | 2.8 | 3.3 | 3.6 | 3 | 3.2 | 3.3 | 3.1 | 2.7409 | 4 | 3.6025 | 4.5 | 2.8 | 3.2 | 3.3 | 4 | 3.8 | 3 | 3.3 |  |  |  |  | 16
| 1 | 4 | 3.8 | 3.4 | < 5 | 3.4 | 3.2 | 3.1 | 3.4 | 3.5 | 3.1 | 3.3 | 3.3 | 2.8 | 2.573 | 4 | 3.7366 | 4.6 | 2.7 | 3.4 | 3.5 | 5 | 3.8 | 3 | 3.3 |  |  |  |  | 3.2952
| 1 | 5 | 3.5 | 3.6 | < 5 | 3.3 | 3.1 | 3.1 | 3.4 | 3.9 | 3.3 | 3.2 | 3.2 | 3.1 | 2.7378 | 4 | 3.9706 | 4.3 | 2.9 | 3.2 | 3.6 | 4 | 3.8 | 3 | 3.5 |  |  |  |  | 191
| 1 | 6 | 3.9 | 3.6 | < 5 | 3.4 | 3.2 | 3.3 | 3.2 | 3.7 | 3.4 | 3.2 | 3.4 | 3 | 2.5112 | 4 | 3.6876 | 4.4 | 2.8 | 3.2 | 3.8 | 4 | 3.9 | 3.1 | 3.5
| Mean |  | 3.7167 | 3.3833 |  | 3.35 | 3.1 | 3.0333 | 3.3333 | 3.6833 | 3.2333 | 3.2333 | 3.35 | 2.9833 | 2.6156 | 4 | 3.7698 | 4.45 | 2.7667 | 3.2833 | 3.5167 | 4.3333 | 3.8333 | 3.0167 | 3.45
| Median |  | 3.7 | 3.4 |  | 3.35 | 3.15 | 3.05 | 3.35 | 3.7 | 3.25 | 3.2 | 3.35 | 3 | 2.5971 | 4 | 3.7446 | 4.45 | 2.75 | 3.2 | 3.55 | 4 | 3.8 | 3 | 3.5
| Std Dev. |  | 0.2137 | 0.22286 |  | 0.054772 | 0.12649 | 0.17512 | 0.08165 | 0.13292 | 0.1633 | 0.05164 | 0.10488 | 0.1169 | 0.10451 | 0 | 0.13133 | 0.13784 | 0.08165 | 0.13292 | 0.19408 | 0.5164 | 0.05164 | 0.075277 | 0.12247
| Rel.Std.Dev. |  | 5.75% | 6.587% |  | 1.635% | 4.08% | 5.773% | 2.449% | 3.609% | 5.05% | 1.597% | 3.131% | 3.919% | 3.996% | 0% | 3.484% | 3.098% | 2.951% | 4.048% | 5.519% | 11.92% | 1.347% | 2.495% | 3.55%
| PDM3 |  | 12.79% | 2.675% |  | 1.663% | -5.924% | -7.947% | 1.157% | 11.78% | -1.877% | -1.877% | 1.663% | -9.464% | -20.62% | 21.39% | 14.4% | 35.05% | -16.04% | -0.36% | 6.721% | 31.5% | 16.33% | -8.453% | 4.698%
| Z-Score (Absolute) |  | 0.86 | 0.11 | 1.87 | 0.04 | 0.52 | 0.67 | 0 | 0.79 | 0.22 | 0.22 | 0.04 | 0.79 | 1.61 | NA | 0.98 | 2.51 | 1.27 | 0.11 | 0.41 | NA | 1.12 | 0.71 | 0.26
Results from laboratories 17 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 555b (Certified Value 6.7 ppm) | [28 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | 5.6 | < 5 | 7 | 6.6 | 7.1 | 4.9385 | 7 | 6.9 | 6.6 | 7.2 | 7.2 | < 10 | < 0.1 | 6.1546 | 7 | 6.8211 | NA | 10 | 6.8 | 7.1 | 7 | 7 | 6.4 |  |  |  |  | 1
| 1 | 2 | 5.2 | < 5 | 6 | 6.5 | 7.3 | 5.0073 | 6.9 | 6.8 | 5.8 | 7.7 | 7.1 | < 10 | < 0.1 | 6.2331 | 7 | 6.8616 | NA | 10 | 5.9 | 6.63 | 7 | 7 | 6.4 |  |  |  |  | 19
| 1 | 3 | 5.2 | < 5 | 6 | 6.2 | 7 | 4.6502 | 7.3 | 6.8 | 6.3 | 7.4 | 7.2 | < 10 | < 0.1 | 6.3797 | 7 | 6.537 | NA | 12 | 7 | 6.11 | 6 | 7 | 6.3 |  |  |  |  | 16
| 1 | 4 | 5.5 | < 5 | 7 | 5.9 | 7.2 | 4.7905 | 7.1 | 6.9 | 6.3 | 6.9 | 7.2 | < 10 | < 0.1 | 6.1878 | 7 | 6.8161 | NA | 10 | 7.7 | 6.59 | 7 | 7 | 6.3 |  |  |  |  | 6.6992
| 1 | 5 | 5.8 | < 5 | 7 | 6.3 | 7.3 | 4.7146 | 7.6 | 7.1 | 6.5 | 7.2 | 7.2 | < 10 | < 0.1 | 6.3215 | 7 | 6.7804 | NA | 10 | 6.6 | 6.32 | 7 | 7 | 6.3 |  |  |  |  | 192
| 1 | 6 | 5.3 | < 5 | 7 | 6.1 | 7.2 | 4.9955 | 6.8 | 6.9 | 6.5 | 7.4 | 7.2 | < 10 | < 0.1 | 6.3826 | 7 | 6.9254 | NA | 10 | 6.8 | 6.56 | 6 | 7 | 6.5
| Mean |  | 5.4333 |  | 6.6667 | 6.2667 | 7.1833 | 4.8494 | 7.1167 | 6.9 | 6.3333 | 7.3 | 7.1833 |  |  | 6.2765 | 7 | 6.7902 |  | 10.3333 | 6.8 | 6.5517 | 6.6667 | 7 | 6.3667
| Median |  | 5.4 |  | 7 | 6.25 | 7.2 | 4.8645 | 7.05 | 6.9 | 6.4 | 7.3 | 7.2 |  |  | 6.2773 | 7 | 6.8186 |  | 10 | 6.8 | 6.575 | 7 | 7 | 6.35
| Std Dev. |  | 0.24221 |  | 0.5164 | 0.2582 | 0.1169 | 0.15198 | 0.29269 | 0.10954 | 0.28752 | 0.26833 | 0.040825 |  |  | 0.098523 | 0 | 0.13358 |  | 0.8165 | 0.5831 | 0.33379 | 0.5164 | 0 | 0.08165
| Rel.Std.Dev. |  | 4.458% |  | 7.746% | 4.12% | 1.627% | 3.134% | 4.113% | 1.588% | 4.54% | 3.676% | 0.5683% |  |  | 1.57% | 0% | 1.967% |  | 7.902% | 8.575% | 5.095% | 7.746% | 0% | 1.282%
| PDM3 |  | -18.9% |  | -0.4851% | -6.456% | 7.227% | -27.61% | 6.232% | 2.998% | -5.461% | 8.969% | 7.227% |  |  | -6.308% | 4.491% | 1.36% |  | 54.25% | 1.505% | -2.202% | -0.4851% | 4.491% | -4.963%
| Z-Score (Absolute) |  | 2.11 | 7.11 | 0 | 0.68 | 0.88 | 3.1 | 0.77 | 0.4 | 0.57 | 1.08 | 0.88 | 2.84 | 11.29 | 0.67 | 0.57 | 0.21 | NA | 6.26 | 0.23 | 0.2 | 0 | 0.57 | 0.51
| Analytical results for Sm in OREAS 555b (Certified Value 3.51 ppm) | [28 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 26 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [21 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | [21 empty] | 2
| 1 | 1 | 3.79 | 4 | 2.64 | 1.7938 | 3.3833 | 4.08 | [21 empty] | 1
| 1 | 2 | 3.565 | 4.1 | 2.76 | 1.7352 | 3.4413 | 4.07 | [21 empty] | 3
| 1 | 3 | 3.553 | 3.6 | 2.78 | 1.7535 | 3.4814 | 3.95 | [21 empty] | 16
| 1 | 4 | 3.553 | 3.7 | 2.77 | 1.8156 | 3.2061 | 4.1 | [21 empty] | 3.5079
| 1 | 5 | 3.323 | 3.8 | 2.7 | 1.7569 | 3.3945 | 4.07 | [21 empty] | 193
| 1 | 6 | 3.685 | 3.6 | 2.75 | 1.6609 | 3.3108 | 4.08
| Mean |  | 3.5782 | 3.8 | 2.7333 | 1.7526 | 3.3696 | 4.0583
| Median |  | 3.559 | 3.75 | 2.755 | 1.7552 | 3.3889 | 4.075
| Std Dev. |  | 0.15674 | 0.20976 | 0.053541 | 0.053662 | 0.098645 | 0.054191
| Rel.Std.Dev. |  | 4.38% | 5.52% | 1.959% | 3.062% | 2.928% | 1.335%
| PDM3 |  | 2.004% | 8.328% | -22.08% | -50.04% | -3.943% | 15.69%
| Z-Score (Absolute) |  | 0.15 | 0.48 | 1.1 | 2.55 | 0.15 | 0.87
| Analytical results for Sn in OREAS 555b (Certified Value 1.54 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 5 | < 10 | 1.58 | 1.48 | 1.5 | 1.4627 | 1.4 | 1.4 | 1.5 | 1.5 | 1.5 | < 10 | 1.29 | 1.3002 | 1.8 | 1.5065 | 1.7 | < 10 | 1.5 | 1.5 | 1.4 | 1.74 | 1.7 |  |  |  |  | 1
| 1 | 2 | < 5 | < 10 | 1.64 | 1.46 | 1.5 | 1.515 | 1.5 | 1.5 | 1.6 | 1.5 | 1.5 | < 10 | 1.29 | 1.1848 | 1.9 | 1.6108 | 1.7 | < 10 | 1.5 | 1.5 | 1.4 | 1.74 | 1.8 |  |  |  |  | 20
| 1 | 3 | < 5 | < 10 | 1.65 | 1.43 | 1.5 | 1.4786 | 1.5 | 1.4 | 1.5 | 1.6 | 1.5 | < 10 | 1.37 | 1.3231 | 1.9 | 1.5967 | 1.7 | < 10 | 1.5 | 1.6 | 1.4 | 1.76 | 1.7 |  |  |  |  | 16
| 1 | 4 | < 5 | < 10 | 1.57 | 1.39 | 1.6 | 1.4899 | 1.4 | 1.4 | 1.7 | 1.6 | 1.5 | < 10 | 1.3 | 1.3309 | 1.9 | 1.5615 | 1.7 | < 10 | 1.5 | 1.6 | 1.4 | 1.67 | 1.7 |  |  |  |  | 1.5356
| 1 | 5 | < 5 | < 10 | 1.73 | 1.39 | 1.6 | 1.4868 | 1.5 | 1.5 | 1.6 | 1.6 | 1.4 | < 10 | 1.3 | 1.2343 | 1.9 | 1.5602 | 1.6 | < 10 | 1.5 | 1.6 | 1.3 | 1.71 | 1.8 |  |  |  |  | 194
| 1 | 6 | < 5 | < 10 | 1.61 | 1.48 | 1.6 | 1.5037 | 1.4 | 1.5 | 1.5 | 1.6 | 1.5 | < 10 | 1.32 | 1.1685 | 1.8 | 1.5848 | 1.6 | < 10 | 1.6 | 1.6 | 1.3 | 1.73 | 1.6
| Mean |  |  |  | 1.63 | 1.4383 | 1.55 | 1.4894 | 1.45 | 1.45 | 1.5667 | 1.5667 | 1.4833 |  | 1.3117 | 1.257 | 1.8667 | 1.5701 | 1.6667 |  | 1.5167 | 1.5667 | 1.3667 | 1.725 | 1.7167
| Median |  |  |  | 1.625 | 1.445 | 1.55 | 1.4884 | 1.45 | 1.45 | 1.55 | 1.6 | 1.5 |  | 1.3 | 1.2673 | 1.9 | 1.5731 | 1.7 |  | 1.5 | 1.6 | 1.4 | 1.735 | 1.7
| Std Dev. |  |  |  | 0.05831 | 0.041673 | 0.054772 | 0.018433 | 0.054772 | 0.054772 | 0.08165 | 0.05164 | 0.040825 |  | 0.030605 | 0.071074 | 0.05164 | 0.036877 | 0.05164 |  | 0.040825 | 0.05164 | 0.05164 | 0.031464 | 0.075277
| Rel.Std.Dev. |  |  |  | 3.577% | 2.897% | 3.534% | 1.238% | 3.777% | 3.777% | 5.212% | 3.296% | 2.752% |  | 2.333% | 5.654% | 2.766% | 2.349% | 3.098% |  | 2.692% | 3.296% | 3.779% | 1.824% | 4.385%
| PDM3 |  |  |  | 6.147% | -6.334% | 0.9374% | -3.007% | -5.575% | -5.575% | 2.023% | 2.023% | -3.404% |  | -14.58% | -18.14% | 21.56% | 2.245% | 8.535% |  | -1.233% | 2.023% | -11% | 12.33% | 11.79%
| Z-Score (Absolute) |  | 5.39 | 19.84 | 0.37 | 0.74 | 0.1 | 0.45 | 0.67 | 0.67 | 0 | 0 | 0.48 | 19.84 | 1.47 | 1.79 | 1.73 | 0.02 | 0.58 | 19.84 | 0.29 | 0 | 1.16 | 0.92 | 0.87
| Analytical results for Sr in OREAS 555b (Certified Value 15.1 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 1
| 1 | 1 | 29.3 | 13.3 | 25.6 | 16.07 | 20.9 | 10.4 | 13.5 | 15.7 | 15.2 | 13.4 | 14.8 | 15.36 | 11 | 14.491 | 22.4 | 16.7761 | 22.7 | 11.3 | 15 | 17.8 | 24.5 | 21 | 15 | 15.2 |  |  |  | 1
| 1 | 2 | 24.8 | 15.1 | 25.1 | 16.58 | 20 | 10.9 | 13.9 | 16.6 | 15 | 13.4 | 15 | 15.06 | 10 | 13.4485 | 20.5 | 17.0753 | 23.5 | 11.3 | 15 | 17 | 29.1 | 21 | 14 | 15.2 |  |  |  | 32
| 1 | 3 | 23.2 | 15.8 | 25.2 | 16.1 | 19.1 | 10.6 | 13.8 | 16.1 | 14.7 | 13.7 | 14.7 | 15.22 | 11 | 14.9336 | 22.8 | 17.0295 | 23.5 | 11.2 | 15 | 16.4 | 28.2 | 20 | 14 | 15.3 |  |  |  | 16
| 1 | 4 | 22.6 | 16.1 | 25.2 | 15.99 | 19 | 11 | 13.9 | 15.8 | 15.4 | 13.9 | 15.3 | 15.28 | 10 | 14.0493 | 21.1 | 16.9081 | 22.9 | 11.2 | 15 | 16.8 | 29 | 21 | 14 | 14.5 |  |  |  | 15.1067
| 1 | 5 | 23.7 | 17.2 | 26.4 | 15.1 | 20.2 | 11.2 | 14.4 | 16.2 | 15.4 | 13.7 | 14.7 | 15.17 | 10 | 16.4306 | 21.3 | 16.3928 | 22.3 | 11.6 | 14 | 17.1 | 17.8 | 21 | 15 | 15.2 |  |  |  | 195
| 1 | 6 | 31.6 | 16.7 | 26.5 | 16.38 | 18.8 | 11.2 | 14.1 | 16.4 | 14.6 | 13.8 | 14.6 | 15.31 | 10 | 16.0849 | 20.9 | 16.7143 | 23.5 | 11.3 | 15 | 17.8 | 18.8 | 21 | 14 | 15
| Mean |  | 25.8667 | 15.7 | 25.6667 | 16.0367 | 19.6667 | 10.8833 | 13.9333 | 16.1333 | 15.05 | 13.65 | 14.85 | 15.2333 | 10.3333 | 14.9063 | 21.5 | 16.816 | 23.0667 | 11.3167 | 14.8333 | 17.15 | 24.5667 | 20.8333 | 14.3333 | 15.0667
| Median |  | 24.25 | 15.95 | 25.4 | 16.085 | 19.55 | 10.95 | 13.9 | 16.15 | 15.1 | 13.7 | 14.75 | 15.25 | 10 | 14.7123 | 21.2 | 16.8421 | 23.2 | 11.3 | 15 | 17.05 | 26.35 | 21 | 14 | 15.2
| Std Dev. |  | 3.6952 | 1.3813 | 0.6314 | 0.50946 | 0.82865 | 0.32506 | 0.30111 | 0.34448 | 0.34496 | 0.20736 | 0.25884 | 0.10801 | 0.5164 | 1.1615 | 0.90111 | 0.24994 | 0.51251 | 0.1472 | 0.40825 | 0.55767 | 5.1469 | 0.40825 | 0.5164 | 0.29439
| Rel.Std.Dev. |  | 14.29% | 8.798% | 2.46% | 3.177% | 4.213% | 2.987% | 2.161% | 2.135% | 2.292% | 1.519% | 1.743% | 0.7091% | 4.997% | 7.792% | 4.191% | 1.486% | 2.222% | 1.301% | 2.752% | 3.252% | 20.95% | 1.96% | 3.603% | 1.954%
| PDM3 |  | 71.23% | 3.928% | 69.9% | 6.156% | 30.19% | -27.96% | -7.767% | 6.796% | -0.375% | -9.642% | -1.699% | 0.8386% | -31.6% | -1.326% | 42.32% | 11.32% | 52.69% | -25.09% | -1.809% | 13.53% | 62.62% | 37.91% | -5.119% | -0.2647%
| Z-Score (Absolute) |  | 4.36 | 0.1 | 4.28 | 0.24 | 1.76 | 1.92 | 0.64 | 0.28 | 0.17 | 0.76 | 0.26 | 0.1 | 2.15 | 0.23 | 2.53 | 0.57 | 3.19 | 1.74 | 0.27 | 0.71 | 3.82 | 2.25 | 0.48 | 0.17
| Analytical results for Ta in OREAS 555b (Certified Value < 0.01 ppm) | [28 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | 23 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [10 empty] | 3
| 1 | 1 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.09 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 1
| 1 | 2 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 5
| 1 | 3 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 16
| 1 | 4 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.06 | < 0.01 | < 0.01 | 0.04 | [10 empty] | < 0.01
| 1 | 5 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 196
| 1 | 6 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.05 | < 0.01 | < 0.01 | < 0.01
| Mean | [14 empty] | 0.068333 |  |  | 0.04
| Median | [14 empty] | 0.07 |  |  | 0.04
| Std Dev. | [14 empty] | 0.013292
| Rel.Std.Dev. | [14 empty] | 19.45%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 555b (Certified Value 0.21 ppm) | [28 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 20 | 26 | [20 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 01g | [20 empty] | 2
| 1 | 1 | 0.214 | 0.2 | 0.15 | 0.17759 | 0.23687 | 0.3 | 0.25 | [20 empty] | 1
| 1 | 2 | 0.209 | 0.2 | 0.16 | 0.17307 | 0.23967 | 0.3 | 0.24 | [20 empty] | 6
| 1 | 3 | 0.209 | 0.2 | 0.17 | 0.17563 | 0.24499 | 0.3 | 0.23 | [20 empty] | 16
| 1 | 4 | 0.211 | 0.2 | 0.16 | 0.18149 | 0.22247 | 0.3 | 0.24 | [20 empty] | 0.20531
| 1 | 5 | 0.207 | 0.2 | 0.15 | 0.17174 | 0.2475 | 0.3 | 0.24 | [20 empty] | 197
| 1 | 6 | 0.21 | 0.2 | 0.16 | 0.16981 | 0.23928 | 0.3 | 0.25
| Mean |  | 0.21 | 0.2 | 0.15833 | 0.17489 | 0.23846 | 0.3 | 0.24167
| Median |  | 0.2095 | 0.2 | 0.16 | 0.17435 | 0.23948 | 0.3 | 0.24
| Std Dev. |  | 0.0023664 | 3.0405e-17 | 0.0075277 | 0.0042533 | 0.0087701 | 0 | 0.0075277
| Rel.Std.Dev. |  | 1.127% | 1.52e-14% | 4.754% | 2.432% | 3.678% | 0% | 3.115%
| PDM3 |  | 2.285% | -2.586% | -22.88% | -14.82% | 16.15% | 46.12% | 17.71%
| Z-Score (Absolute) |  | 0 | NA | 1.1 | 0.75 | 0.61 | NA | 0.67
| Analytical results for Te in OREAS 555b (Certified Value 0.054 ppm) | [28 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 | 26 | 28 | [5 empty] | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.03 | < 5 | < 0.1 | 0.04 | 0.04 | 0.08499 | 0.04 | 0.06 | 0.05 | 0.08 | 0.04 | < 10 | 0.04 | < 0.05 | < 0.05 | < 5 | NA | < 10 | < 0.1 | 0.04 | 0.05 | 0.05 | [5 empty] | 1
| 1 | 2 | 0.04 | < 5 | < 0.1 | 0.05 | 0.07 | < 0.04 | 0.05 | 0.07 | 0.06 | 0.06 | 0.06 | < 10 | 0.06 | < 0.05 | < 0.05 | < 5 | NA | < 10 | < 0.1 | 0.05 | 0.08 | 0.06 | [5 empty] | 7
| 1 | 3 | 0.07 | < 5 | < 0.1 | 0.05 | 0.04 | 0.04393 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 | < 10 | 0.06 | < 0.05 | < 0.05 | < 5 | NA | < 10 | < 0.1 | 0.04 | < 0.05 | 0.05 | [5 empty] | 16
| 1 | 4 | 0.04 | < 5 | < 0.1 | 0.07 | 0.06 | 0.05557 | 0.04 | 0.07 | 0.05 | 0.05 | 0.04 | < 10 | 0.04 | < 0.05 | < 0.05 | < 5 | NA | < 10 | < 0.1 | 0.04 | < 0.05 | 0.06 | [5 empty] | 0.053516
| 1 | 5 | 0.08 | < 5 | < 0.1 | 0.06 | 0.1 | 0.05804 | 0.03 | 0.06 | 0.06 | 0.04 | 0.04 | < 10 | 0.06 | < 0.05 | < 0.05 | < 5 | NA | < 10 | < 0.1 | 0.04 | 0.06 | 0.06 | [5 empty] | 198
| 1 | 6 | 0.03 | < 5 | < 0.1 | 0.06 | 0.07 | 0.06848 | 0.06 | 0.05 | 0.05 | 0.05 | 0.04 | < 10 | 0.05 | < 0.05 | < 0.05 | < 5 | NA | < 10 | < 0.1 | 0.04 | 0.08 | 0.06
| Mean |  | 0.048333 |  |  | 0.055 | 0.063333 | 0.062202 | 0.045 | 0.061667 | 0.051667 | 0.053333 | 0.045 |  | 0.051667 | [6 empty] | 0.041667 | 0.0675 | 0.056667
| Median |  | 0.04 |  |  | 0.055 | 0.065 | 0.05804 | 0.045 | 0.06 | 0.05 | 0.05 | 0.04 |  | 0.055 | [6 empty] | 0.04 | 0.07 | 0.06
| Std Dev. |  | 0.02137 |  |  | 0.010488 | 0.022509 | 0.015442 | 0.010488 | 0.0075277 | 0.0075277 | 0.015055 | 0.0083666 |  | 0.0098319 | [6 empty] | 0.0040825 | 0.015 | 0.005164
| Rel.Std.Dev. |  | 44.21% |  |  | 19.07% | 35.54% | 24.83% | 23.31% | 12.21% | 14.57% | 28.23% | 18.59% |  | 19.03% | [6 empty] | 9.798% | 22.22% | 9.113%
| PDM3 |  | -9.684% |  |  | 2.774% | 18.35% | 16.23% | -15.91% | 15.23% | -3.455% | -0.3405% | -15.91% |  | -3.455% | [6 empty] | -22.14% | 26.13% | 5.888%
| Z-Score (Absolute) |  | 0.4 | 197.98 | 0.27 | 0.13 | 0.81 | 0.15 | 0.67 | 0.67 | 0.13 | 0 | 0.67 | 400.27 | 0.13 | 2.29 | 2.29 | 197.98 | NA | 400.27 | 0.27 | 0.94 | 0 | 0.27
| Analytical results for Th in OREAS 555b (Certified Value 6.32 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | [5 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | [5 empty] | 2
| 1 | 1 | 12.9 | 5.5 | 6.82 | 7.7 | 5.1 | 5.4 | 7 | 6.1 | 5.5 | 5.9 | 5.31 | 6.3554 | 4.7379 | 8.15 | 6.0984 | < 20 | 5.5 | 6.7 | 9.82 | 8.15 | 6.4 | 6.7 | [5 empty] | 1
| 1 | 2 | 11 | 5.5 | 6.74 | 7.4 | 5.3 | 5.5 | 7 | 6 | 5.5 | 5.9 | 5.53 | 6.9116 | 4.7232 | 7.87 | 6.6453 | < 20 | 5.6 | 6.73 | 10.4 | 8.33 | 6.3 | 6.8 | [5 empty] | 8
| 1 | 3 | 9.8 | 6.3 | 6.54 | 7.2 | 5.2 | 5.6 | 7.3 | 5.8 | 5.4 | 5.9 | 5.7 | 6.4776 | 4.8209 | 8.19 | 6.083 | < 20 | 5.8 | 6.78 | 10.1 | 7.88 | 6.3 | 6.7 | [5 empty] | 16
| 1 | 4 | 10.4 | 6.2 | 6.61 | 7 | 5.2 | 5.5 | 7.1 | 5.7 | 5.5 | 5.9 | 5.43 | 6.1643 | 4.9679 | 7.91 | 5.9107 | < 20 | 5.8 | 7 | 10.3 | 8.08 | 6.5 | 6.4 | [5 empty] | 6.3228
| 1 | 5 | 10.4 | 6.6 | 6.66 | 7.4 | 5.2 | 5.6 | 7.3 | 6.4 | 5.4 | 5.7 | 5.37 | 6.8273 | 4.676 | 7.94 | 6.4622 | < 20 | 5.8 | 7.01 | 8.09 | 8 | 6.4 | 6.7 | [5 empty] | 199
| 1 | 6 | 13.4 | 6.3 | 6.78 | 7.2 | 5.4 | 5.5 | 7.5 | 6.2 | 5.5 | 5.6 | 5.49 | 6.6938 | 4.4947 | 8.04 | 6.1575 | < 20 | 5.9 | 7.11 | 8.04 | 8.08 | 6.4 | 6.5
| Mean |  | 11.3167 | 6.0667 | 6.6917 | 7.3167 | 5.2333 | 5.5167 | 7.2 | 6.0333 | 5.4667 | 5.8167 | 5.4717 | 6.5717 | 4.7368 | 8.0167 | 6.2262 |  | 5.7333 | 6.8883 | 9.4583 | 8.0867 | 6.3833 | 6.6333
| Median |  | 10.7 | 6.25 | 6.7 | 7.3 | 5.2 | 5.5 | 7.2 | 6.05 | 5.5 | 5.9 | 5.46 | 6.5857 | 4.7305 | 7.99 | 6.1279 |  | 5.8 | 6.89 | 9.96 | 8.08 | 6.4 | 6.7
| Std Dev. |  | 1.4784 | 0.45898 | 0.10704 | 0.24014 | 0.10328 | 0.075277 | 0.2 | 0.2582 | 0.05164 | 0.13292 | 0.13717 | 0.28892 | 0.15683 | 0.13201 | 0.27292 |  | 0.15055 | 0.17244 | 1.0974 | 0.15069 | 0.075277 | 0.15055
| Rel.Std.Dev. |  | 13.06% | 7.566% | 1.6% | 3.282% | 1.973% | 1.365% | 2.778% | 4.28% | 0.9446% | 2.285% | 2.507% | 4.396% | 3.311% | 1.647% | 4.383% |  | 2.626% | 2.503% | 11.6% | 1.863% | 1.179% | 2.27%
| PDM3 |  | 78.98% | -4.051% | 5.834% | 15.72% | -17.23% | -12.75% | 13.87% | -4.578% | -13.54% | -8.005% | -13.46% | 3.936% | -25.08% | 26.79% | -1.528% |  | -9.323% | 8.945% | 49.59% | 27.9% | 0.9576% | 4.912%
| Z-Score (Absolute) |  | 4.12 | 0.35 | 0.18 | 0.71 | 1.06 | 0.82 | 0.62 | 0.38 | 0.86 | 0.56 | 0.86 | 0.08 | 1.48 | 1.31 | 0.21 | 3 | 0.63 | 0.35 | 2.54 | 1.37 | 0.08 | 0.13
| Analytical results for Ti in OREAS 555b (Indicative Value 0.002 wt.%) | [28 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 27 | 28 | [5 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | [5 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | [5 empty] | 3
| 1 | 1 | < 0.01 | 0.002 | < 0.01 | 0.0007 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0011868 | 0.0017 | 0.0012514 | < 0.01 | < 0.01 | < 0.005 | 0.0015 | 0.0015 | 0.001 | < 0.005 | [5 empty] | 1
| 1 | 2 | < 0.01 | 0.002 | < 0.01 | 0.0008 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0010027 | 0.0017 | 0.0013401 | < 0.01 | < 0.01 | < 0.005 | 0.0028 | 0.0014 | 0.001 | < 0.005 | [5 empty] | 25
| 1 | 3 | < 0.01 | 0.002 | < 0.01 | 0.0007 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0011675 | 0.0017 | 0.0013393 | < 0.01 | < 0.01 | < 0.005 | 0.0028 | 0.0014 | 0.001 | < 0.005 | [5 empty] | 16
| 1 | 4 | < 0.01 | 0.003 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0010518 | 0.0017 | 0.0012724 | < 0.01 | < 0.01 | < 0.005 | 0.0028 | 0.0012 | 0.001 | < 0.005 | [5 empty] | 0.0015157
| 1 | 5 | < 0.01 | 0.003 | < 0.01 | 0.0005 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0016477 | 0.0018 | 0.0013309 | < 0.01 | < 0.01 | < 0.005 | 0.0013 | 0.0015 | 0.002 | < 0.005 | [5 empty] | 31
| 1 | 6 | < 0.01 | 0.002 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0015609 | 0.0018 | 0.0012629 | < 0.01 | < 0.01 | < 0.005 | < 0.001 | 0.0016 | 0.001 | < 0.005
| Mean |  |  | 0.0023333 |  | 0.00065 | [8 empty] | 0.0012696 | 0.0017333 | 0.0012995 |  |  |  | 0.00224 | 0.0014333 | 0.0011667
| Median |  |  | 0.002 |  | 0.00065 | [8 empty] | 0.0011772 | 0.0017 | 0.0013016 |  |  |  | 0.0028 | 0.00145 | 0.001
| Std Dev. |  |  | 0.0005164 |  | 0.00010488 | [8 empty] | 0.00026974 | 5.164e-05 | 4.1469e-05 |  |  |  | 0.00077006 | 0.00013663 | 0.00040825
| Rel.Std.Dev. |  |  | 22.13% |  | 16.14% | [8 empty] | 21.25% | 2.979% | 3.191% |  |  |  | 34.38% | 9.532% | 34.99%
| PDM3 |  |  | 53.94% |  | -57.12% | [8 empty] | -16.24% | 14.36% | -14.27% |  |  |  | 47.79% | -5.435% | -23.03%
| Z-Score (Absolute) |  | 1.84 | 0.12 | 1.84 | 1.36 | 1.84 | 0 | 0 | 0 | 0 | 0 | 0 | 1.4 | 0.91 | 0.56 | 0.88 | 1.84 | 1.84 | 0 | 0.4 | 0.78 | 0.98 | 0
| Analytical results for Tl in OREAS 555b (Indicative Value 0.36 ppm) | [28 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 2
| 1 | 1 | 29.6 | 0.56 | < 10 | 0.68 | 0.04 | 0.49 | 0.32337 | 0.08 | 0.05 | 0.1 | 0.14 | 0.09 | 0.38787 | 0.61 | 0.25474 | 0.311 | 0.61 | 2 | 0.06 | 0.45 | 0.75 | 0.74 | < 0.5 | 0.56 |  |  |  | 1
| 1 | 2 | 27.8 | 0.54 | < 10 | 0.69 | 0.05 | 0.51 | 0.29573 | 0.08 | 0.06 | 0.09 | 0.14 | 0.09 | 0.38255 | 0.6 | 0.26924 | 0.353 | 0.62 | 2 | 0.06 | 0.44 | 0.71 | 0.76 | < 0.5 | 0.54 |  |  |  | 26
| 1 | 3 | 30.4 | 0.56 | < 10 | 0.68 | 0.04 | 0.49 | 0.29341 | 0.09 | 0.06 | 0.1 | 0.13 | 0.09 | 0.42443 | 0.62 | 0.26078 | 0.334 | 0.62 | 2 | 0.05 | 0.45 | 0.72 | 0.75 | < 0.5 | 0.53 |  |  |  | 16
| 1 | 4 | 26.3 | 0.53 | < 10 | 0.68 | 0.04 | 0.5 | 0.21919 | 0.08 | 0.05 | 0.09 | 0.15 | 0.09 | 0.31015 | 0.6 | 0.23606 | 0.335 | 0.61 | 2 | 0.05 | 0.43 | 0.74 | 0.74 | < 0.5 | 0.5 |  |  |  | 0.35805
| 1 | 5 | 27.7 | 0.56 | < 10 | 0.67 | 0.04 | 0.5 | 0.29207 | 0.08 | 0.05 | 0.08 | 0.12 | 0.08 | 0.3669 | 0.6 | 0.28119 | 0.332 | 0.59 | 2 | 0.05 | 0.41 | 0.53 | 0.74 | < 0.5 | 0.51 |  |  |  | 32
| 1 | 6 | 29.8 | 0.55 | < 10 | 0.68 | 0.05 | 0.51 | 0.30316 | 0.08 | 0.05 | 0.07 | 0.12 | 0.09 | 0.32957 | 0.6 | 0.26879 | 0.299 | 0.59 | < 2 | 0.06 | 0.43 | 0.63 | 0.74 | < 0.5 | 0.53
| Mean |  | 28.6 | 0.55 |  | 0.68 | 0.043333 | 0.5 | 0.28782 | 0.081667 | 0.053333 | 0.088333 | 0.13333 | 0.088333 | 0.36691 | 0.605 | 0.2618 | 0.32733 | 0.60667 | 2 | 0.055 | 0.435 | 0.68 | 0.745 |  | 0.52833
| Median |  | 28.7 | 0.555 |  | 0.68 | 0.04 | 0.5 | 0.29457 | 0.08 | 0.05 | 0.09 | 0.135 | 0.09 | 0.37472 | 0.6 | 0.26479 | 0.333 | 0.61 | 2 | 0.055 | 0.435 | 0.715 | 0.74 |  | 0.53
| Std Dev. |  | 1.5761 | 0.012649 |  | 0.0063246 | 0.005164 | 0.0089443 | 0.035556 | 0.0040825 | 0.005164 | 0.01169 | 0.012111 | 0.0040825 | 0.041495 | 0.0083666 | 0.015452 | 0.019253 | 0.013663 | 0 | 0.0054772 | 0.015166 | 0.084853 | 0.0083666 |  | 0.02137
| Rel.Std.Dev. |  | 5.511% | 2.3% |  | 0.9301% | 11.92% | 1.789% | 12.35% | 4.999% | 9.682% | 13.23% | 9.083% | 4.622% | 11.31% | 1.383% | 5.902% | 5.882% | 2.252% | 0% | 9.959% | 3.486% | 12.48% | 1.123% |  | 4.045%
| PDM3 |  | 7888% | 53.61% |  | 89.92% | -87.9% | 39.65% | -19.61% | -77.19% | -85.1% | -75.33% | -62.76% | -75.33% | 2.476% | 68.97% | -26.88% | -8.578% | 69.44% | 458.6% | -84.64% | 21.49% | 89.92% | 108.1% |  | 47.56%
| Z-Score (Absolute) |  | 79.41 | 0.51 | 13.03 | 0.88 | 0.91 | 0.37 | 0.22 | 0.8 | 0.88 | 0.78 | 0.66 | 0.78 | 0 | 0.67 | 0.3 | 0.11 | 0.67 | NA | 0.88 | 0.19 | 0.88 | 1.06 | 0.33 | 0.45
| Analytical results for Tm in OREAS 555b (Indicative Value 0.051 ppm) | [28 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 04 | 05 | 18 | 26 | [23 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | [23 empty] | 3
|  |  | 01g | 0.5g | 0.2g | 01g | [23 empty] | 3
| 1 | 1 | 0.056 | < 0.1 | < 0.05 | 0.07 | [23 empty] | 1
| 1 | 2 | 0.058 | < 0.1 | 0.053453 | 0.07 | [23 empty] | 10
| 1 | 3 | 0.055 | < 0.1 | 0.050004 | 0.06 | [23 empty] | 16
| 1 | 4 | 0.054 | < 0.1 | < 0.05 | 0.07 | [23 empty] | 0.050722
| 1 | 5 | 0.05 | < 0.1 | < 0.05 | 0.07 | [23 empty] | 30
| 1 | 6 | 0.06 | < 0.1 | < 0.05 | 0.06
| Mean |  | 0.0555 |  | 0.051729 | 0.066667
| Median |  | 0.0555 |  | 0.051729 | 0.07
| Std Dev. |  | 0.0034496 |  | 0.0024389 | 0.005164
| Rel.Std.Dev. |  | 6.216% |  | 4.715% | 7.746%
| PDM3 |  | 9.419% |  | 1.984% | 31.43%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 1.52 | 1.13
| Analytical results for U in OREAS 555b (Certified Value 3.49 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | 3.07 | 3.61 | 3.8 | 3.18 | 3.2 | 3.62 | 3.19 | 3.29 | 3.37 | 3.12 | 3.5977 | 2.3105 | 3.92 | 3.2638 | 3.23 | 3.53 | < 10 | 3 | 3.31 | 4 | 3.97 | 3.5 | 3.92 |  |  |  |  | 1
| 1 | 2 | 3.06 | 3.52 | 3.7 | 3.31 | 3.31 | 3.55 | 3.13 | 3.31 | 3.29 | 3.25 | 3.5436 | 2.2707 | 3.92 | 3.4417 | 3.09 | 3.51 | < 10 | 3 | 3.5 | 4.13 | 4.03 | 3.5 | 3.75 |  |  |  |  | 11
| 1 | 3 | 3.5 | 3.53 | 3.5 | 3.25 | 3.34 | 3.67 | 3.24 | 3.22 | 3.29 | 3.33 | 3.7563 | 2.4681 | 3.8 | 3.201 | 3.17 | 3.62 | < 10 | 3 | 3.61 | 4.18 | 3.88 | 3.5 | 3.75 |  |  |  |  | 16
| 1 | 4 | 3.41 | 3.56 | 3.4 | 3.26 | 3.29 | 3.52 | 3.14 | 3.35 | 3.41 | 3.22 | 3.5736 | 2.4202 | 3.97 | 3.0704 | 3.24 | 3.6 | < 10 | 3.1 | 3.61 | 4.11 | 3.92 | 3.6 | 3.68 |  |  |  |  | 3.491
| 1 | 5 | 3.74 | 3.46 | 3.6 | 3.24 | 3.31 | 3.77 | 3.32 | 3.3 | 3.26 | 3.2 | 3.7497 | 2.3184 | 4 | 3.5202 | 3.23 | 3.55 | < 10 | 3.1 | 3.69 | 3.63 | 3.97 | 3.6 | 3.85 |  |  |  |  | 200
| 1 | 6 | 3.4 | 3.54 | 3.5 | 3.41 | 3.33 | 3.68 | 3.25 | 3.34 | 3.23 | 3.34 | 3.78 | 2.2688 | 3.84 | 3.1746 | 3.17 | 3.56 | < 10 | 3.2 | 3.67 | 3.75 | 3.96 | 3.6 | 3.72
| Mean |  | 3.3633 | 3.5367 | 3.5833 | 3.275 | 3.2967 | 3.635 | 3.2117 | 3.3017 | 3.3083 | 3.2433 | 3.6668 | 2.3428 | 3.9083 | 3.2786 | 3.1883 | 3.5617 |  | 3.0667 | 3.565 | 3.9667 | 3.955 | 3.55 | 3.7783
| Median |  | 3.405 | 3.535 | 3.55 | 3.255 | 3.31 | 3.645 | 3.215 | 3.305 | 3.29 | 3.235 | 3.6737 | 2.3144 | 3.92 | 3.2324 | 3.2 | 3.555 |  | 3.05 | 3.61 | 4.055 | 3.965 | 3.55 | 3.75
| Std Dev. |  | 0.26158 | 0.049261 | 0.1472 | 0.078166 | 0.050465 | 0.091815 | 0.072503 | 0.046224 | 0.06824 | 0.083106 | 0.10613 | 0.082472 | 0.076004 | 0.17052 | 0.057417 | 0.041673 |  | 0.08165 | 0.14139 | 0.22545 | 0.050892 | 0.054772 | 0.089312
| Rel.Std.Dev. |  | 7.778% | 1.393% | 4.108% | 2.387% | 1.531% | 2.526% | 2.257% | 1.4% | 2.063% | 2.562% | 2.894% | 3.52% | 1.945% | 5.201% | 1.801% | 1.17% |  | 2.662% | 3.966% | 5.684% | 1.287% | 1.543% | 2.364%
| PDM3 |  | -3.657% | 1.308% | 2.645% | -6.187% | -5.566% | 4.125% | -8.001% | -5.423% | -5.232% | -7.094% | 5.037% | -32.89% | 11.95% | -6.084% | -8.67% | 2.025% |  | -12.15% | 2.12% | 13.63% | 13.29% | 1.69% | 8.231%
| Z-Score (Absolute) |  | 0.48 | 0 | 0.13 | 0.73 | 0.67 | 0.27 | 0.91 | 0.66 | 0.64 | 0.82 | 0.36 | 3.33 | 1.04 | 0.72 | 0.97 | 0.07 | 4.08 | 1.31 | 0.08 | 1.2 | 1.17 | 0.04 | 0.67
| Analytical results for V in OREAS 555b (Certified Value 62 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 0
| 1 | 1 | 68 | 59 | 78.8 | 68 | 82 | 47.2 | 54 | 65 | 57 | 56 | 58 | 59.94 | 48 | 68.7774 | 98 | 57.7163 | 117 | 50 | 62 | 69 | 151 | 98 | 63 | 62 |  |  |  | 1
| 1 | 2 | 58.2 | 66 | 78.2 | 69 | 81 | 48.8 | 56 | 67 | 57 | 56 | 58 | 59.19 | 49.2 | 63.0166 | 94 | 60.6831 | 119 | 50 | 60 | 71 | 167 | 98 | 60 | 61 |  |  |  | 38
| 1 | 3 | 59.8 | 70 | 79.2 | 71 | 77 | 47.6 | 57 | 67 | 56 | 57 | 58 | 59.46 | 50.8 | 66.3733 | 98 | 59.0662 | 113 | 50 | 59 | 68 | 161 | 95 | 61 | 61 |  |  |  | 16
| 1 | 4 | 60.5 | 72 | 78.3 | 70 | 79 | 49.5 | 57 | 67 | 57 | 57 | 59 | 59.83 | 47.9 | 64.8 | 93 | 59.5661 | 119 | 49 | 61 | 69 | 167 | 100 | 63 | 60 |  |  |  | 62.0724
| 1 | 5 | 58.6 | 77 | 79.5 | 69 | 81 | 49.6 | 58 | 66 | 58 | 56 | 57 | 59.72 | 48 | 72.1303 | 90 | 59.9146 | 116 | 50 | 59 | 70 | 131 | 96 | 64 | 62 |  |  |  | 201
| 1 | 6 | 68 | 74 | 79.6 | 69 | 78 | 50.8 | 56 | 66 | 58 | 58 | 56 | 59.68 | 49.7 | 70.2171 | 90 | 58.8061 | 119 | 50 | 59 | 70 | 136 | 98 | 63 | 63
| Mean |  | 62.1833 | 69.6667 | 78.9333 | 69.3333 | 79.6667 | 48.9167 | 56.3333 | 66.3333 | 57.1667 | 56.6667 | 57.6667 | 59.6367 | 48.9333 | 67.5525 | 93.8333 | 59.2921 | 117.1667 | 49.8333 | 60 | 69.5 | 152.1667 | 97.5 | 62.3333 | 61.5
| Median |  | 60.15 | 71 | 79 | 69 | 80 | 49.15 | 56.5 | 66.5 | 57 | 56.5 | 58 | 59.7 | 48.6 | 67.5754 | 93.5 | 59.3162 | 118 | 50 | 59.5 | 69.5 | 156 | 98 | 63 | 61.5
| Std Dev. |  | 4.5801 | 6.4083 | 0.59889 | 1.0328 | 1.9664 | 1.3452 | 1.3663 | 0.8165 | 0.75277 | 0.8165 | 1.0328 | 0.27164 | 1.1793 | 3.438 | 3.6009 | 1.0166 | 2.4014 | 0.40825 | 1.2649 | 1.0488 | 15.6769 | 1.7607 | 1.5055 | 1.0488
| Rel.Std.Dev. |  | 7.366% | 9.199% | 0.7587% | 1.49% | 2.468% | 2.75% | 2.425% | 1.231% | 1.317% | 1.441% | 1.791% | 0.4555% | 2.41% | 5.089% | 3.838% | 1.715% | 2.05% | 0.8192% | 2.108% | 1.509% | 10.3% | 1.806% | 2.415% | 1.705%
| PDM3 |  | 0.1787% | 12.23% | 27.16% | 11.7% | 28.34% | -21.19% | -9.246% | 6.864% | -7.903% | -8.709% | -7.098% | -3.924% | -21.17% | 8.829% | 51.17% | -4.479% | 88.76% | -19.72% | -3.339% | 11.97% | 145.1% | 57.07% | 0.4204% | -0.9221%
| Z-Score (Absolute) |  | 0.01 | 0.77 | 1.73 | 0.73 | 1.81 | 1.38 | 0.61 | 0.42 | 0.53 | 0.58 | 0.48 | 0.27 | 1.38 | 0.55 | 3.28 | 0.31 | 5.7 | 1.29 | 0.23 | 0.75 | 9.33 | 3.66 | 0.01 | 0.08
| Analytical results for W in OREAS 555b (Certified Value 1.08 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 5 | 0.86 | < 5 | 1.04 | 0.9 | 1.25 | 1.05 | 0.97 | 1.3 | 1.18 | 0.97 | 1.2 | 0.90513 | 1.1 | 1.0545 | 1.2 | < 10 | 0.9 | 1.08 | 1.35 | 2 | 1.2 | 1.19 |  |  |  |  | 1
| 1 | 2 | < 5 | 0.85 | < 5 | 1.08 | 0.9 | 1.3 | 1.12 | 1.02 | 1.31 | 1.16 | 0.98 | 1.3 | 0.94301 | 1 | 1.076 | 1.2 | < 10 | 0.9 | 1.13 | 1.13 | < 2 | 1.2 | 1.23 |  |  |  |  | 12
| 1 | 3 | < 5 | 0.77 | < 5 | 1.04 | 0.9 | 1.27 | 1.1 | 1.04 | 1.32 | 1.12 | 0.94 | 1.3 | 1.1295 | 1 | 1.2137 | 1.3 | < 10 | 0.9 | 1.07 | 1.16 | 2 | 1.2 | 1.21 |  |  |  |  | 16
| 1 | 4 | < 5 | 0.72 | < 5 | 1 | 0.8 | 1.29 | 1.06 | 0.99 | 1.28 | 1.13 | 0.94 | 1.4 | 1.1175 | 1 | < 1 | 1.1 | < 10 | 0.9 | 1.02 | 1.11 | < 2 | 1.2 | 1.15 |  |  |  |  | 1.0849
| 1 | 5 | < 5 | 0.81 | < 5 | 0.9 | 0.9 | 1.3 | 1.14 | 1.02 | 1.32 | 1.13 | 0.99 | 1.3 | 1.1789 | 1 | 1.1181 | 1.2 | < 10 | 0.9 | 1.11 | 1.04 | < 2 | 1.1 | 1.19 |  |  |  |  | 202
| 1 | 6 | < 5 | 0.81 | < 5 | 1.01 | 0.9 | 1.33 | 1.04 | 1.01 | 1.28 | 1.08 | 0.94 | 1.3 | 1.13 | 1 | 1.0772 | 1.2 | < 10 | 1 | 1.07 | 1.07 | < 2 | 1.2 | 1.16
| Mean |  |  | 0.80333 |  | 1.0117 | 0.88333 | 1.29 | 1.085 | 1.0083 | 1.3017 | 1.1333 | 0.96 | 1.3 | 1.0674 | 1.0167 | 1.1079 | 1.2 |  | 0.91667 | 1.08 | 1.1433 | 2 | 1.1833 | 1.1883
| Median |  |  | 0.81 |  | 1.025 | 0.9 | 1.295 | 1.08 | 1.015 | 1.305 | 1.13 | 0.955 | 1.3 | 1.1235 | 1 | 1.0772 | 1.2 |  | 0.9 | 1.075 | 1.12 | 2 | 1.2 | 1.19
| Std Dev. |  |  | 0.052026 |  | 0.061455 | 0.040825 | 0.027568 | 0.040866 | 0.024833 | 0.018348 | 0.034448 | 0.022804 | 0.063246 | 0.11361 | 0.040825 | 0.063458 | 0.063246 |  | 0.040825 | 0.037947 | 0.10985 | 0 | 0.040825 | 0.029944
| Rel.Std.Dev. |  |  | 6.476% |  | 6.075% | 4.622% | 2.137% | 3.766% | 2.463% | 1.41% | 3.04% | 2.375% | 4.865% | 10.64% | 4.016% | 5.728% | 5.27% |  | 4.454% | 3.514% | 9.608% | 0% | 3.45% | 2.52%
| PDM3 |  |  | -25.95% |  | -6.747% | -18.58% | 18.91% | 0.01226% | -7.055% | 19.98% | 4.467% | -11.51% | 19.83% | -1.614% | -6.287% | 2.123% | 10.61% |  | -15.5% | -0.4486% | 5.389% | 84.35% | 9.076% | 9.537%
| Z-Score (Absolute) |  | 7.27 | 1.76 | 7.27 | 0.65 | 1.33 | 0.83 | 0.26 | 0.66 | 0.9 | 0 | 0.92 | 0.89 | 0.35 | 0.62 | 0.67 | 0.35 | 20.57 | 1.15 | 0.28 | 0.05 | 1.06 | 0.27 | 0.29
| Analytical results for Y in OREAS 555b (Certified Value 3.94 ppm) | [28 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [6 empty] | 2
| 1 | 1 | < 5 | 4.13 | 4.21 | 3.1 | 3.64 | 4.37 | 3.57 | 3.42 | 3.61 | < 10 | 3.68 | 4.1213 | 4.75 | 3.8558 | 4.8 | 3.2 | 3.8 | 3.84 | 5.3 | 4.76 | 3.88 | [6 empty] | 1
| 1 | 2 | < 5 | 4.16 | 4.26 | 3.3 | 3.75 | 4.42 | 3.47 | 3.4 | 3.63 | < 10 | 3.76 | 3.8614 | 4.85 | 3.885 | 4.9 | 3.2 | 3.8 | 3.6 | 5.7 | 4.85 | 3.75 | [6 empty] | 13
| 1 | 3 | < 5 | 4.14 | 4.09 | 3.21 | 3.75 | 4.31 | 3.52 | 3.54 | 3.56 | < 10 | 3.89 | 4.223 | 4.85 | 3.9196 | 5.3 | 3.2 | 3.8 | 3.56 | 5.65 | 4.59 | 3.8 | [6 empty] | 16
| 1 | 4 | < 5 | 4.2 | 3.87 | 3.31 | 3.79 | 4.3 | 3.57 | 3.48 | 3.69 | < 10 | 3.63 | 3.9751 | 4.87 | 3.8178 | 4.8 | 3.2 | 3.8 | 3.69 | 5.51 | 4.7 | 3.66 | [6 empty] | 3.9366
| 1 | 5 | < 5 | 4.17 | 4.06 | 3.32 | 3.82 | 4.31 | 3.63 | 3.53 | 3.54 | < 10 | 3.71 | 4.4865 | 4.84 | 3.7682 | 4.8 | 3.2 | 3.7 | 3.78 | 4.22 | 4.67 | 3.96 | [6 empty] | 203
| 1 | 6 | < 5 | 4.2 | 3.88 | 3.44 | 3.8 | 4.35 | 3.49 | 3.43 | 3.53 | < 10 | 3.79 | 4.4332 | 4.84 | 3.8336 | 4.9 | 3.2 | 3.6 | 3.84 | 4.33 | 4.77 | 3.8
| Mean |  |  | 4.1667 | 4.0617 | 3.28 | 3.7583 | 4.3433 | 3.5417 | 3.4667 | 3.5933 |  | 3.7433 | 4.1834 | 4.8333 | 3.8467 | 4.9167 | 3.2 | 3.75 | 3.7183 | 5.1183 | 4.7233 | 3.8083
| Median |  |  | 4.165 | 4.075 | 3.305 | 3.77 | 4.33 | 3.545 | 3.455 | 3.585 |  | 3.735 | 4.1721 | 4.845 | 3.8447 | 4.85 | 3.2 | 3.8 | 3.735 | 5.405 | 4.73 | 3.8
| Std Dev. |  |  | 0.029439 | 0.16241 | 0.11472 | 0.064317 | 0.046332 | 0.05947 | 0.059217 | 0.061536 |  | 0.091579 | 0.24769 | 0.042269 | 0.053 | 0.19408 | 4.8648e-16 | 0.083666 | 0.12106 | 0.66871 | 0.090259 | 0.10362
| Rel.Std.Dev. |  |  | 0.7065% | 3.999% | 3.497% | 1.711% | 1.067% | 1.679% | 1.708% | 1.713% |  | 2.446% | 5.921% | 0.8745% | 1.378% | 3.947% | 1.52e-14% | 2.231% | 3.256% | 13.07% | 1.911% | 2.721%
| PDM3 |  |  | 5.845% | 3.178% | -16.68% | -4.528% | 10.33% | -10.03% | -11.94% | -8.719% |  | -4.909% | 6.27% | 22.78% | -2.284% | 24.9% | -18.71% | -4.74% | -5.544% | 30.02% | 19.99% | -3.258%
| Z-Score (Absolute) |  | 2.46 | 0.67 | 0.48 | 0.99 | 0.09 | 1.01 | 0.5 | 0.64 | 0.4 | 2.24 | 0.12 | 0.71 | 1.93 | 0.07 | 2.09 | 1.14 | 0.11 | 0.17 | 2.47 | 1.72 | 0
| Analytical results for Yb in OREAS 555b (Certified Value 0.32 ppm) | [28 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 04 | 05 | 14 | 16 | 18 | 20 | 23 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.25g | 01g | [19 empty] | 2
| 1 | 1 | 0.372 | 0.4 | 0.27 | 0.19118 | 0.32698 | 0.5 | 0.4 | 0.42 | [19 empty] | 1
| 1 | 2 | 0.363 | 0.4 | 0.28 | 0.20268 | 0.33799 | 0.5 | 0.4 | 0.43 | [19 empty] | 14
| 1 | 3 | 0.372 | 0.4 | 0.29 | 0.2024 | 0.32325 | 0.5 | 0.4 | 0.4 | [19 empty] | 16
| 1 | 4 | 0.367 | 0.4 | 0.27 | 0.20044 | 0.3112 | 0.5 | 0.4 | 0.4 | [19 empty] | 0.31535
| 1 | 5 | 0.357 | 0.4 | 0.28 | 0.20044 | 0.34583 | 0.5 | 0.4 | 0.39 | [19 empty] | 204
| 1 | 6 | 0.369 | 0.4 | 0.28 | 0.19034 | 0.32772 | 0.5 | 0.4 | 0.39
| Mean |  | 0.36667 | 0.4 | 0.27833 | 0.19791 | 0.32883 | 0.5 | 0.4 | 0.405
| Median |  | 0.368 | 0.4 | 0.28 | 0.20044 | 0.32735 | 0.5 | 0.4 | 0.4
| Std Dev. |  | 0.0058195 | 6.0809e-17 | 0.0075277 | 0.0056242 | 0.011996 | 0 | 6.0809e-17 | 0.016432
| Rel.Std.Dev. |  | 1.587% | 1.52e-14% | 2.705% | 2.842% | 3.648% | 0% | 1.52e-14% | 4.057%
| PDM3 |  | 16.27% | 26.84% | -11.74% | -37.24% | 4.275% | 58.55% | 26.84% | 28.43%
| Z-Score (Absolute) |  | 0.51 | NA | 0.67 | 1.75 | 0 | NA | NA | 1.02
| Analytical results for Zn in OREAS 555b (Certified Value 26.6 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 2
| 1 | 1 | 26.7 | 28.3 | 27.4 | 27 | 25.2 | 26.1 | 27.9022 | 24 | 26 | 26 | 26 | 26 | 29.32 | 31 | 27.6454 | 22.6 | 28 | 26.0255 | 26 | 29.5 | 26.8 | 28 | 23.6 | 27 | 27 | 27 |  | 1
| 1 | 2 | 27.4 | 29.3 | 26.6 | 26 | 23.9 | 29.3 | 26.2857 | 25 | 27 | 26 | 25 | 25 | 28.94 | 32.6 | 27.6936 | 23.42 | 29 | 26.3412 | 26 | 29.6 | 26.4 | 28 | 22.7 | 27 | 27 | 25 |  | 15
| 1 | 3 | 25.6 | 32.7 | 27.6 | 27 | 22.2 | 31.9 | 25.91 | 25 | 26 | 27 | 24 | 26 | 29.18 | 35 | 28.2273 | 21.65 | 29 | 26.4842 | 26 | 29.7 | 27.2 | 27 | 22.7 | 27 | 27 | 24 |  | 16
| 1 | 4 | 26.8 | 29.6 | 26.9 | 27 | 23.4 | 29 | 27.054 | 25 | 26 | 25 | 25 | 27 | 29.16 | 35.1 | 27.3648 | 22.71 | 28 | 25.9268 | 26 | 30.5 | 25.8 | 27 | 21.5 | 26 | 27 | 25 |  | 26.6028
| 1 | 5 | 26.2 | 32 | 27.2 | 27 | 24.8 | 38.3 | 26.3115 | 25 | 27 | 26 | 25 | 26 | 29.32 | 32.5 | 28.2999 | 22.81 | 29 | 26.2238 | 27 | 30.6 | 26.6 | 26 | 24.3 | 28 | 27 | 27 |  | 205
| 1 | 6 | 26.6 | 29.8 | 27.6 | 26 | 24.6 | 28.6 | 24.8733 | 25 | 26 | 26 | 25 | 26 | 29.18 | 32.7 | 27.7348 | 22.4 | 27 | 26.1701 | 26 | 29.9 | 27.2 | 27 | 23.4 | 27 | 27 | 25
| Mean |  | 26.55 | 30.2833 | 27.2167 | 26.6667 | 24.0167 | 30.5333 | 26.3894 | 24.8333 | 26.3333 | 26 | 25 | 26 | 29.1833 | 33.15 | 27.8276 | 22.5983 | 28.3333 | 26.1953 | 26.1667 | 29.9667 | 26.6667 | 27.1667 | 23.0333 | 27 | 27 | 25.5
| Median |  | 26.65 | 29.7 | 27.3 | 27 | 24.25 | 29.15 | 26.2986 | 25 | 26 | 26 | 25 | 26 | 29.18 | 32.65 | 27.7142 | 22.655 | 28.5 | 26.1969 | 26 | 29.8 | 26.7 | 27 | 23.05 | 27 | 27 | 25
| Std Dev. |  | 0.60581 | 1.6964 | 0.40208 | 0.5164 | 1.0998 | 4.2297 | 1.0272 | 0.40825 | 0.5164 | 0.63246 | 0.63246 | 0.63246 | 0.13938 | 1.5984 | 0.36241 | 0.57811 | 0.8165 | 0.20353 | 0.40825 | 0.47188 | 0.53166 | 0.75277 | 0.96264 | 0.63246 | 0 | 1.2247
| Rel.Std.Dev. |  | 2.282% | 5.602% | 1.477% | 1.936% | 4.58% | 13.85% | 3.892% | 1.644% | 1.961% | 2.433% | 2.53% | 2.433% | 0.4776% | 4.822% | 1.302% | 2.558% | 2.882% | 0.777% | 1.56% | 1.575% | 1.994% | 2.771% | 4.179% | 2.342% | 0% | 4.803%
| PDM3 |  | -0.1985% | 13.84% | 2.307% | 0.24% | -9.721% | 14.77% | -0.8021% | -6.651% | -1.013% | -2.266% | -6.025% | -2.266% | 9.7% | 24.61% | 4.604% | -15.05% | 6.505% | -1.532% | -1.639% | 12.64% | 0.24% | 2.12% | -13.42% | 1.493% | 1.493% | -4.145%
| Z-Score (Absolute) |  | 0.05 | 2.89 | 0.48 | 0.05 | 2.04 | 3.08 | 0.17 | 1.39 | 0.22 | 0.48 | 1.26 | 0.48 | 2.02 | 5.14 | 0.96 | 3.15 | 1.36 | 0.32 | 0.35 | 2.64 | 0.05 | 0.44 | 2.81 | 0.31 | 0.31 | 0.87
| Analytical results for Zr in OREAS 555b (Certified Value 10.1 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [5 empty] | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [5 empty] | 1
| 1 | 1 | 15 | 11.4 | 11.7 | 11.5 | 8.6 | 8.4 | 9.2 | 9 | 9.7 | 9 | 9.51 | 8.68 | 10.9103 | 12.2 | 9.1725 | 13.3 | 6.3 | 8.9 | 9.1 | 10 | 10.9 | 10 | [5 empty] | 1
| 1 | 2 | 14.2 | 11 | 11.6 | 11.1 | 8.8 | 8.6 | 9.3 | 9.1 | 9.6 | 9.1 | 9.62 | 8.25 | 13.1007 | 12.1 | 9.3449 | 12.4 | 6.2 | 8.7 | 10.4 | 11.5 | 10.7 | 10 | [5 empty] | 18
| 1 | 3 | 11.6 | 11.4 | 11.7 | 11.2 | 8.6 | 8.7 | 9.3 | 8.6 | 9.8 | 8.9 | 9.61 | 9.33 | 12.2767 | 11.6 | 9.4347 | 13.3 | 6.3 | 8.8 | 9.6 | 11.3 | 10.5 | 9.9 | [5 empty] | 16
| 1 | 4 | 13.5 | 11.2 | 11.8 | 10.5 | 8.8 | 8.7 | 9.4 | 8.9 | 9.9 | 9 | 9.63 | 8.64 | 13.0645 | 11.6 | 9.2095 | 12.7 | 6.3 | 9 | 9.8 | 11.3 | 10.9 | 9.5 | [5 empty] | 10.1145
| 1 | 5 | 12.6 | 11.6 | 11.1 | 10.9 | 8.9 | 8.9 | 9.2 | 9.5 | 9.7 | 8.7 | 9.61 | 8.65 | 11.7212 | 11.5 | 9.1349 | 12.6 | 6.3 | 9.1 | 10.3 | 7.7 | 10.5 | 10 | [5 empty] | 206
| 1 | 6 | 15.1 | 11.7 | 11.9 | 10.6 | 9.2 | 8.8 | 9 | 11.6 | 9.7 | 8.8 | 9.55 | 8.9 | 11.5653 | 11.7 | 9.2638 | 12.9 | 6.3 | 8.7 | 10.1 | 9.3 | 10.7 | 9.5
| Mean |  | 13.6667 | 11.3833 | 11.6333 | 10.9667 | 8.8167 | 8.6833 | 9.2333 | 9.45 | 9.7333 | 8.9167 | 9.5883 | 8.7417 | 12.1064 | 11.7833 | 9.2601 | 12.8667 | 6.2833 | 8.8667 | 9.8833 | 10.1833 | 10.7 | 9.8167
| Median |  | 13.85 | 11.4 | 11.7 | 11 | 8.8 | 8.7 | 9.25 | 9.05 | 9.7 | 8.95 | 9.61 | 8.665 | 11.9989 | 11.65 | 9.2367 | 12.8 | 6.3 | 8.85 | 9.95 | 10.65 | 10.7 | 9.95
| Std Dev. |  | 1.3823 | 0.25626 | 0.28048 | 0.37771 | 0.22286 | 0.17224 | 0.13663 | 1.0932 | 0.10328 | 0.1472 | 0.047504 | 0.35651 | 0.87262 | 0.29269 | 0.11278 | 0.37238 | 0.040825 | 0.1633 | 0.48751 | 1.4972 | 0.17889 | 0.24833
| Rel.Std.Dev. |  | 10.11% | 2.251% | 2.411% | 3.444% | 2.528% | 1.984% | 1.48% | 11.57% | 1.061% | 1.651% | 0.4954% | 4.078% | 7.208% | 2.484% | 1.218% | 2.894% | 0.6497% | 1.842% | 4.933% | 14.7% | 1.672% | 2.53%
| PDM3 |  | 35.12% | 12.54% | 15.02% | 8.425% | -12.83% | -14.15% | -8.712% | -6.57% | -3.768% | -11.84% | -5.202% | -13.57% | 19.69% | 16.5% | -8.447% | 27.21% | -37.88% | -12.34% | -2.285% | 0.6806% | 5.789% | -2.945%
| Z-Score (Absolute) |  | 2.79 | 1.15 | 1.33 | 0.85 | 0.69 | 0.78 | 0.39 | 0.23 | 0.03 | 0.61 | 0.13 | 0.74 | 1.67 | 1.44 | 0.37 | 2.21 | 2.5 | 0.65 | 0.08 | 0.29 | 0.66 | 0.03
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C240: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C258: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C277: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C295: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C332: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C350: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C368: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C387: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C424: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C443: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C462: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C480: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C499: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C535: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C553: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C571: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C589: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C644: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C681: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C717: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C735: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C754: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C772: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C790: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C808: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C826: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C844: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C862: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1026: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1116: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1188: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1206: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 555b (Indicative Value 0.383 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.38 | 1
| 1 | 3 | 0.39 | 16
| 1 | 4 | 0.39 | 0.38333
| 1 | 5 | 0.4 | 34
| 1 | 6 | 0.37
| Mean |  | 0.38333
| Median |  | 0.385
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 3.159%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 555b (Indicative Value 12.17 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 12.2 | 1
| 1 | 2 | 12 | 29
| 1 | 3 | 12.2 | 16
| 1 | 4 | 12.4 | 12.1667
| 1 | 5 | 12.3 | 35
| 1 | 6 | 11.9
| Mean |  | 12.1667
| Median |  | 12.2
| Std Dev. |  | 0.18619
| Rel.Std.Dev. |  | 1.53%
| PDM3 |  | -2.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 555b (Indicative Value 247 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 248 | 1
| 1 | 2 | 244 | 30
| 1 | 3 | 245 | 16
| 1 | 4 | 254 | 247.1667
| 1 | 5 | 249 | 36
| 1 | 6 | 243
| Mean |  | 247.1667
| Median |  | 246.5
| Std Dev. |  | 4.0702
| Rel.Std.Dev. |  | 1.647%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 555b (Indicative Value 5.85 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.9 | 1
| 1 | 2 | 5.8 | 3
| 1 | 3 | 6.1 | 16
| 1 | 4 | 5.6 | 5.85
| 1 | 5 | 6.1 | 37
| 1 | 6 | 5.6
| Mean |  | 5.85
| Median |  | 5.85
| Std Dev. |  | 0.22583
| Rel.Std.Dev. |  | 3.86%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 555b (Indicative Value 1.95 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.9 | 1
| 1 | 2 | 1.9 | 4
| 1 | 3 | 1.9 | 16
| 1 | 4 | 1.9 | 1.95
| 1 | 5 | 1.9 | 38
| 1 | 6 | 2.2
| Mean |  | 1.95
| Median |  | 1.9
| Std Dev. |  | 0.12247
| Rel.Std.Dev. |  | 6.281%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 555b (Indicative Value 0.19 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.19 | 1
| 1 | 2 | 0.19 | 33
| 1 | 3 | 0.19 | 16
| 1 | 4 | 0.19 | 0.19
| 1 | 5 | 0.19 | 39
| 1 | 6 | 0.19
| Mean |  | 0.19
| Median |  | 0.19
| Std Dev. |  | 3.0405e-17
| Rel.Std.Dev. |  | 1.6e-14%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 555b (Indicative Value 0.58 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 5
| 1 | 3 | 0.6 | 16
| 1 | 4 | 0.6 | 0.58333
| 1 | 5 | 0.6 | 40
| 1 | 6 | 0.5
| Mean |  | 0.58333
| Median |  | 0.6
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 6.999%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 555b (Indicative Value 142 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 143 | 1
| 1 | 2 | 140 | 6
| 1 | 3 | 140 | 16
| 1 | 4 | 146 | 142
| 1 | 5 | 142 | 41
| 1 | 6 | 141
| Mean |  | 142
| Median |  | 141.5
| Std Dev. |  | 2.2804
| Rel.Std.Dev. |  | 1.606%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 555b (Indicative Value 1.17 wt.%) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.15 | 1
| 1 | 2 | 1.17 | 36
| 1 | 3 | 1.2 | 16
| 1 | 4 | 1.2 | 1.17
| 1 | 5 | 1.2 | 42
| 1 | 6 | 1.1
| Mean |  | 1.17
| Median |  | 1.185
| Std Dev. |  | 0.04
| Rel.Std.Dev. |  | 3.419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 555b (Indicative Value 104 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 104 | 1
| 1 | 2 | 104 | 37
| 1 | 3 | 104 | 16
| 1 | 4 | 106 | 104.1667
| 1 | 5 | 104 | 43
| 1 | 6 | 103
| Mean |  | 104.1667
| Median |  | 104
| Std Dev. |  | 0.98319
| Rel.Std.Dev. |  | 0.9439%
| PDM3 |  | -3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 555b (Indicative Value 2.15 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2 | 9
| 1 | 3 | 2 | 16
| 1 | 4 | 2 | 2.15
| 1 | 5 | 2.8 | 44
| 1 | 6 | 2
| Mean |  | 2.15
| Median |  | 2
| Std Dev. |  | 0.32094
| Rel.Std.Dev. |  | 14.93%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 555b (Indicative Value 4.84 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.83 | 1
| 1 | 2 | 4.88 | 39
| 1 | 3 | 4.84 | 16
| 1 | 4 | 4.93 | 4.8383
| 1 | 5 | 4.88 | 45
| 1 | 6 | 4.67
| Mean |  | 4.8383
| Median |  | 4.86
| Std Dev. |  | 0.089759
| Rel.Std.Dev. |  | 1.855%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 555b (Indicative Value 4.83 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 4.86 | 10
| 1 | 3 | 4.71 | 16
| 1 | 4 | 4.95 | 4.8267
| 1 | 5 | 4.93 | 46
| 1 | 6 | 4.71
| Mean |  | 4.8267
| Median |  | 4.83
| Std Dev. |  | 0.10482
| Rel.Std.Dev. |  | 2.172%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 555b (Indicative Value 2.9 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.96 | 1
| 1 | 2 | 2.84 | 11
| 1 | 3 | 2.86 | 16
| 1 | 4 | 2.96 | 2.8983
| 1 | 5 | 2.94 | 47
| 1 | 6 | 2.83
| Mean |  | 2.8983
| Median |  | 2.9
| Std Dev. |  | 0.061455
| Rel.Std.Dev. |  | 2.12%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 555b (Indicative Value 0.96 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.96 | 1
| 1 | 2 | 1.01 | 12
| 1 | 3 | 0.96 | 16
| 1 | 4 | 0.97 | 0.95833
| 1 | 5 | 0.92 | 48
| 1 | 6 | 0.93
| Mean |  | 0.95833
| Median |  | 0.96
| Std Dev. |  | 0.031885
| Rel.Std.Dev. |  | 3.327%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 555b (Indicative Value 5.41 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 5.42 | 1
| 1 | 2 | 5.32 | 43
| 1 | 3 | 5.41 | 16
| 1 | 4 | 5.54 | 5.4083
| 1 | 5 | 5.44 | 49
| 1 | 6 | 5.32
| Mean |  | 5.4083
| Median |  | 5.415
| Std Dev. |  | 0.082563
| Rel.Std.Dev. |  | 1.527%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 555b (Indicative Value 28.3 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 27.4 | 1
| 1 | 2 | 28.4 | 13
| 1 | 3 | 29 | 16
| 1 | 4 | 28.5 | 28.3
| 1 | 5 | 28.8 | 50
| 1 | 6 | 27.7
| Mean |  | 28.3
| Median |  | 28.45
| Std Dev. |  | 0.6261
| Rel.Std.Dev. |  | 2.212%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 555b (Indicative Value 5.63 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.6 | 1
| 1 | 2 | 5.5 | 14
| 1 | 3 | 5.7 | 16
| 1 | 4 | 5.8 | 5.6333
| 1 | 5 | 5.5 | 34
| 1 | 6 | 5.7
| Mean |  | 5.6333
| Median |  | 5.65
| Std Dev. |  | 0.12111
| Rel.Std.Dev. |  | 2.15%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 555b (Indicative Value 5.97 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6 | 1
| 1 | 2 | 5.9 | 19
| 1 | 3 | 5.8 | 16
| 1 | 4 | 6.1 | 5.9667
| 1 | 5 | 6.1 | 35
| 1 | 6 | 5.9
| Mean |  | 5.9667
| Median |  | 5.95
| Std Dev. |  | 0.12111
| Rel.Std.Dev. |  | 2.03%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 555b (Indicative Value 0.93 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.91 | 1
| 1 | 2 | 0.9 | 3
| 1 | 3 | 0.92 | 16
| 1 | 4 | 0.96 | 0.92667
| 1 | 5 | 0.95 | 36
| 1 | 6 | 0.92
| Mean |  | 0.92667
| Median |  | 0.92
| Std Dev. |  | 0.023381
| Rel.Std.Dev. |  | 2.523%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 555b (Indicative Value 3.1 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.11 | 1
| 1 | 2 | 3.04 | 31
| 1 | 3 | 3.1 | 16
| 1 | 4 | 3.17 | 3.095
| 1 | 5 | 3.12 | 37
| 1 | 6 | 3.03
| Mean |  | 3.095
| Median |  | 3.105
| Std Dev. |  | 0.05244
| Rel.Std.Dev. |  | 1.694%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 555b (Indicative Value 79 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 78 | 1
| 1 | 2 | 78 | 5
| 1 | 3 | 78 | 16
| 1 | 4 | 81 | 78.6667
| 1 | 5 | 80 | 38
| 1 | 6 | 77
| Mean |  | 78.6667
| Median |  | 78
| Std Dev. |  | 1.5055
| Rel.Std.Dev. |  | 1.914%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 555b (Indicative Value 62 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 62 | 1
| 1 | 2 | 60 | 33
| 1 | 3 | 62 | 16
| 1 | 4 | 63 | 61.5
| 1 | 5 | 62 | 39
| 1 | 6 | 60
| Mean |  | 61.5
| Median |  | 62
| Std Dev. |  | 1.2247
| Rel.Std.Dev. |  | 1.991%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 555b (Indicative Value 5.12 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 5.14 | 1
| 1 | 2 | 5.03 | 34
| 1 | 3 | 5.14 | 16
| 1 | 4 | 5.22 | 5.12
| 1 | 5 | 5.16 | 40
| 1 | 6 | 5.03
| Mean |  | 5.12
| Median |  | 5.14
| Std Dev. |  | 0.075631
| Rel.Std.Dev. |  | 1.477%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 555b (Indicative Value 0.115 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.114 | 1
| 1 | 2 | 0.118 | 35
| 1 | 3 | 0.12 | 16
| 1 | 4 | 0.113 | 0.1145
| 1 | 5 | 0.113 | 41
| 1 | 6 | 0.109
| Mean |  | 0.1145
| Median |  | 0.1135
| Std Dev. |  | 0.003937
| Rel.Std.Dev. |  | 3.438%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 555b (Indicative Value 19.6 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 19.8 | 1
| 1 | 2 | 19.5 | 7
| 1 | 3 | 19.6 | 16
| 1 | 4 | 20.2 | 19.5667
| 1 | 5 | 19.5 | 42
| 1 | 6 | 18.8
| Mean |  | 19.5667
| Median |  | 19.55
| Std Dev. |  | 0.45898
| Rel.Std.Dev. |  | 2.346%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 555b (Indicative Value 0.073 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.07 | 8
| 1 | 3 | 0.08 | 16
| 1 | 4 | 0.08 | 0.073333
| 1 | 5 | 0.07 | 43
| 1 | 6 | 0.07
| Mean |  | 0.073333
| Median |  | 0.07
| Std Dev. |  | 0.005164
| Rel.Std.Dev. |  | 7.042%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 555b (Indicative Value 23.4 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 23.2 | 1
| 1 | 2 | 23 | 9
| 1 | 3 | 23.5 | 16
| 1 | 4 | 23.7 | 23.35
| 1 | 5 | 23.6 | 44
| 1 | 6 | 23.1
| Mean |  | 23.35
| Median |  | 23.35
| Std Dev. |  | 0.2881
| Rel.Std.Dev. |  | 1.234%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 555b (Indicative Value 65 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 64.9 | 1
| 1 | 2 | 65.1 | 10
| 1 | 3 | 64.9 | 16
| 1 | 4 | 64.7 | 64.5667
| 1 | 5 | 64.2 | 45
| 1 | 6 | 63.6
| Mean |  | 64.5667
| Median |  | 64.8
| Std Dev. |  | 0.56451
| Rel.Std.Dev. |  | 0.8743%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 555b (Indicative Value 254 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 248 | 1
| 1 | 2 | 258 | 40
| 1 | 3 | 265 | 16
| 1 | 4 | 259 | 253.6667
| 1 | 5 | 255 | 46
| 1 | 6 | 237
| Mean |  | 253.6667
| Median |  | 256.5
| Std Dev. |  | 9.8725
| Rel.Std.Dev. |  | 3.892%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 555b (Indicative Value < 0.002 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | < 0.002 | 1
| 1 | 2 | < 0.002 | 41
| 1 | 3 | < 0.002 | 16
| 1 | 4 | < 0.002 | < 0.002
| 1 | 5 | < 0.002 | 47
| 1 | 6 | < 0.002
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 555b (Indicative Value 15.8 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 15.5 | 1
| 1 | 2 | 15.7 | 11
| 1 | 3 | 15.8 | 16
| 1 | 4 | 15.8 | 15.8167
| 1 | 5 | 16 | 48
| 1 | 6 | 16.1
| Mean |  | 15.8167
| Median |  | 15.8
| Std Dev. |  | 0.2137
| Rel.Std.Dev. |  | 1.351%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 555b (Indicative Value 17.4 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 17.6 | 1
| 1 | 2 | 17.4 | 12
| 1 | 3 | 17.3 | 16
| 1 | 4 | 17.5 | 17.3833
| 1 | 5 | 17.6 | 49
| 1 | 6 | 16.9
| Mean |  | 17.3833
| Median |  | 17.45
| Std Dev. |  | 0.26394
| Rel.Std.Dev. |  | 1.518%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 555b (Indicative Value 103 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 95.4 | 1
| 1 | 2 | 95.2 | 13
| 1 | 3 | 95.9 | 16
| 1 | 4 | 94.3 | 102.5167
| 1 | 5 | 142 | 50
| 1 | 6 | 92.3
| Mean |  | 102.5167
| Median |  | 95.3
| Std Dev. |  | 19.3845
| Rel.Std.Dev. |  | 18.91%
| PDM3 |  | -3.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 555b (Indicative Value 4.19 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 4.14 | 1
| 1 | 2 | 4.14 | 28
| 1 | 3 | 4.12 | 16
| 1 | 4 | 4.28 | 4.1867
| 1 | 5 | 4.31 | 34
| 1 | 6 | 4.13
| Mean |  | 4.1867
| Median |  | 4.14
| Std Dev. |  | 0.084774
| Rel.Std.Dev. |  | 2.025%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 555b (Indicative Value 13 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13 | 1
| 1 | 2 | 12.6 | 1
| 1 | 3 | 12.8 | 16
| 1 | 4 | 13.4 | 12.95
| 1 | 5 | 13.3 | 35
| 1 | 6 | 12.6
| Mean |  | 12.95
| Median |  | 12.9
| Std Dev. |  | 0.34496
| Rel.Std.Dev. |  | 2.664%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 555b (Indicative Value 8.12 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 7.9 | 1
| 1 | 2 | 8.2 | 3
| 1 | 3 | 8.2 | 16
| 1 | 4 | 8.3 | 8.1167
| 1 | 5 | 7.9 | 36
| 1 | 6 | 8.2
| Mean |  | 8.1167
| Median |  | 8.2
| Std Dev. |  | 0.17224
| Rel.Std.Dev. |  | 2.122%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 555b (Indicative Value 2.97 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.92 | 1
| 1 | 2 | 2.92 | 20
| 1 | 3 | 3.12 | 16
| 1 | 4 | 3.03 | 2.97
| 1 | 5 | 2.94 | 37
| 1 | 6 | 2.89
| Mean |  | 2.97
| Median |  | 2.93
| Std Dev. |  | 0.087636
| Rel.Std.Dev. |  | 2.951%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 555b (Indicative Value 72 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 72 | 1
| 1 | 2 | 71 | 32
| 1 | 3 | 71 | 16
| 1 | 4 | 73 | 71.5
| 1 | 5 | 72 | 38
| 1 | 6 | 70
| Mean |  | 71.5
| Median |  | 71.5
| Std Dev. |  | 1.0488
| Rel.Std.Dev. |  | 1.467%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 555b (Indicative Value 0.35 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.34 | 5
| 1 | 3 | 0.35 | 16
| 1 | 4 | 0.37 | 0.35333
| 1 | 5 | 0.36 | 39
| 1 | 6 | 0.34
| Mean |  | 0.35333
| Median |  | 0.355
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 3.428%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 555b (Indicative Value 0.74 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.74 | 1
| 1 | 2 | 0.71 | 6
| 1 | 3 | 0.74 | 16
| 1 | 4 | 0.77 | 0.73667
| 1 | 5 | 0.72 | 40
| 1 | 6 | 0.74
| Mean |  | 0.73667
| Median |  | 0.74
| Std Dev. |  | 0.020656
| Rel.Std.Dev. |  | 2.804%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 555b (Indicative Value 17.6 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 17.6 | 1
| 1 | 2 | 17.5 | 8
| 1 | 3 | 17.6 | 16
| 1 | 4 | 17.6 | 17.6
| 1 | 5 | 17.8 | 41
| 1 | 6 | 17.5
| Mean |  | 17.6
| Median |  | 17.6
| Std Dev. |  | 0.10954
| Rel.Std.Dev. |  | 0.6224%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 555b (Indicative Value 0.782 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.791 | 1
| 1 | 2 | 0.77 | 36
| 1 | 3 | 0.789 | 16
| 1 | 4 | 0.797 | 0.78233
| 1 | 5 | 0.784 | 42
| 1 | 6 | 0.763
| Mean |  | 0.78233
| Median |  | 0.7865
| Std Dev. |  | 0.01314
| Rel.Std.Dev. |  | 1.68%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 555b (Indicative Value 10.5 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 9.78 | 1
| 1 | 2 | 9.68 | 11
| 1 | 3 | 10.3 | 16
| 1 | 4 | 13.3 | 10.4867
| 1 | 5 | 10 | 43
| 1 | 6 | 9.86
| Mean |  | 10.4867
| Median |  | 9.93
| Std Dev. |  | 1.3949
| Rel.Std.Dev. |  | 13.3%
| PDM3 |  | -3.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 555b (Indicative Value 374 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 374 | 1
| 1 | 2 | 370 | 38
| 1 | 3 | 371 | 16
| 1 | 4 | 382 | 374.1667
| 1 | 5 | 378 | 44
| 1 | 6 | 370
| Mean |  | 374.1667
| Median |  | 372.5
| Std Dev. |  | 4.916
| Rel.Std.Dev. |  | 1.314%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 555b (Indicative Value 3.7 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.6 | 12
| 1 | 3 | 3.6 | 16
| 1 | 4 | 3.6 | 3.7
| 1 | 5 | 4.5 | 45
| 1 | 6 | 3.5
| Mean |  | 3.7
| Median |  | 3.6
| Std Dev. |  | 0.4
| Rel.Std.Dev. |  | 10.81%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 555b (Indicative Value 27.7 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 27.4 | 1
| 1 | 2 | 27.7 | 13
| 1 | 3 | 27.6 | 16
| 1 | 4 | 28.3 | 27.7333
| 1 | 5 | 28.2 | 46
| 1 | 6 | 27.2
| Mean |  | 27.7333
| Median |  | 27.65
| Std Dev. |  | 0.43665
| Rel.Std.Dev. |  | 1.574%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 555b (Indicative Value 2.9 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.87 | 1
| 1 | 2 | 2.88 | 14
| 1 | 3 | 2.86 | 16
| 1 | 4 | 3.03 | 2.8967
| 1 | 5 | 2.91 | 47
| 1 | 6 | 2.83
| Mean |  | 2.8967
| Median |  | 2.875
| Std Dev. |  | 0.070333
| Rel.Std.Dev. |  | 2.428%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 555b (Indicative Value 67 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 65 | 1
| 1 | 2 | 65 | 15
| 1 | 3 | 71 | 16
| 1 | 4 | 66 | 66.8333
| 1 | 5 | 68 | 48
| 1 | 6 | 66
| Mean |  | 66.8333
| Median |  | 66
| Std Dev. |  | 2.3166
| Rel.Std.Dev. |  | 3.466%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 555b (Indicative Value 198 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 194 | 1
| 1 | 2 | 196 | 18
| 1 | 3 | 205 | 16
| 1 | 4 | 197 | 197.6667
| 1 | 5 | 192 | 49
| 1 | 6 | 202
| Mean |  | 197.6667
| Median |  | 196.5
| Std Dev. |  | 4.9261
| Rel.Std.Dev. |  | 2.492%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.