## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 906b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 52.3638 | 50.6713 | 54.0562 | 50.6918 | 54.0358
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppb) | 51.8384 | 49.5919 | 54.0848 | 50.0251 | 53.6516
Sequential Leach
Cu-Res(4A), Copper residual (post leach) (wt.%) | 0.038743 | 0.036722 | 0.040764 | 0.037591 | 0.039894
Cu-Sol(CN), Copper soluble (cyanidation) (wt.%) | 0.0154 | 0.013667 | 0.017133 | 0.014603 | 0.016197
Cu-Sol(H2SO4), Copper soluble (sulphuric) (wt.%) | 0.25975 | 0.25208 | 0.26742 | 0.25648 | 0.26302
Cu-Tot(Calc), Copper total (calculated) (wt.%) | 0.31258 | 0.30383 | 0.32134 | 0.30803 | 0.31714
Gas / Liquid Pycnometry
SG, Specific Gravity (Unity) | 2.7413 | 2.7109 | 2.7716 | 2.7237 | 2.7589
4-Acid Digestion
Ag, Silver (ppm) | 0.70937 | 0.66148 | 0.75726 | 0.6882 | 0.73054
Al, Aluminium (wt.%) | 7.4554 | 7.2438 | 7.667 | 7.3318 | 7.579
As, Arsenic (ppm) | 13.6874 | 12.5816 | 14.7932 | 13.1951 | 14.1798
Ba, Barium (ppm) | 2819.2182 | 2733.6348 | 2904.8015 | 2755.0047 | 2883.4316
Be, Beryllium (ppm) | 2.986 | 2.8125 | 3.1595 | 2.8919 | 3.0801
Bi, Bismuth (ppm) | 9.993 | 9.605 | 10.381 | 9.7458 | 10.2402
Ca, Calcium (wt.%) | 0.64419 | 0.62722 | 0.66115 | 0.633 | 0.65538
Cd, Cadmium (ppm) | 0.41753 | 0.3719 | 0.46316 | 0.3827 | 0.45237
Ce, Cerium (ppm) | 92.1932 | 88.1101 | 96.2763 | 89.6138 | 94.7726
Co, Cobalt (ppm) | 17.5817 | 16.9145 | 18.2488 | 17.1246 | 18.0388
Cr, Chromium (ppm) | 10.1781 | 8.712 | 11.6443 | 9.3526 | 11.0037
Cs, Caesium (ppm) | 6.97 | 6.7305 | 7.2096 | 6.8172 | 7.1229
Cu, Copper (wt.%) | 0.31434 | 0.30695 | 0.32173 | 0.3088 | 0.31989
Dy, Dysprosium (ppm) | 3.705 | 3.3856 | 4.0244 | 3.5593 | 3.8508
Er, Erbium (ppm) | 1.0734 | 0.9335 | 1.2134 | 1.0036 | 1.1433
Eu, Europium (ppm) | 1.5052 | 1.2941 | 1.7164 | 1.4314 | 1.5791
Fe, Iron (wt.%) | 4.8529 | 4.7556 | 4.9502 | 4.7892 | 4.9166
Ga, Gallium (ppm) | 28.1329 | 27.1838 | 29.0819 | 27.361 | 28.9048
Gd, Gadolinium (ppm) | 5.9026 | 5.2459 | 6.5593 | 5.5722 | 6.233
Ge, Germanium (ppm) | 0.20724 | 0.15424 | 0.26024 | 0.19617 | 0.21831
Hf, Hafnium (ppm) | 7.1516 | 6.7891 | 7.5141 | 6.9264 | 7.3767
Ho, Holmium (ppm) | 0.50291 | 0.43426 | 0.57157 | 0.47602 | 0.52981
In, Indium (ppm) | 1.1595 | 1.0984 | 1.2206 | 1.1288 | 1.1902
K, Potassium (wt.%) | 2.8848 | 2.8126 | 2.9569 | 2.8198 | 2.9497
La, Lanthanum (ppm) | 44.9691 | 42.9229 | 47.0152 | 43.7463 | 46.1918
Li, Lithium (ppm) | 20.5939 | 19.6221 | 21.5657 | 20.0485 | 21.1393
Lu, Lutetium (ppm) | 0.084998 | 0.074831 | 0.095165 | IND | IND
Mg, Magnesium (wt.%) | 0.23304 | 0.22238 | 0.24371 | 0.22781 | 0.23828
Mn, Manganese (wt.%) | 0.032119 | 0.030899 | 0.033338 | 0.031331 | 0.032906
Mo, Molybdenum (ppm) | 4.9125 | 4.5681 | 5.2569 | 4.7357 | 5.0893
Na, Sodium (wt.%) | 2.6259 | 2.5488 | 2.703 | 2.5778 | 2.674
Nb, Niobium (ppm) | 18.2642 | 17.4582 | 19.0702 | 17.6015 | 18.9269
Nd, Neodymium (ppm) | 38.6178 | 35.0815 | 42.154 | 37.6728 | 39.5627
Ni, Nickel (ppm) | 3.2247 | 2.8308 | 3.6187 | 3.0054 | 3.4441
P, Phosphorus (wt.%) | 0.028074 | 0.027021 | 0.029127 | 0.02683 | 0.029318
Pb, Lead (ppm) | 31.6988 | 30.222 | 33.1756 | 30.604 | 32.7936
Pr, Praseodymium (ppm) | 10.1972 | 9.3927 | 11.0016 | 9.8306 | 10.5637
Rb, Rubidium (ppm) | 138.48 | 133.3216 | 143.6383 | 135.0719 | 141.888
Re, Rhenium (ppm) | 0.0022146 | 0.0012391 | 0.0031901 | IND | IND
S, Sulphur (wt.%) | 0.044263 | 0.041727 | 0.046799 | 0.042716 | 0.04581
Sb, Antimony (ppm) | 2.1562 | 2.029 | 2.2833 | 2.0359 | 2.2764
Sc, Scandium (ppm) | 4.7135 | 4.3381 | 5.0889 | 4.5773 | 4.8496
Se, Selenium (ppm) | 5.1088 | 4.3147 | 5.9029 | 4.8797 | 5.3379
Sm, Samarium (ppm) | 7.3708 | 6.8781 | 7.8635 | 7.0169 | 7.7247
Sn, Tin (ppm) | 4.2394 | 3.9611 | 4.5178 | 4.0525 | 4.4264
Sr, Strontium (ppm) | 164.7968 | 159.5117 | 170.0819 | 161.4689 | 168.1247
Ta, Tantalum (ppm) | 1.3654 | 1.2686 | 1.4621 | 1.3036 | 1.4271
Tb, Terbium (ppm) | 0.79156 | 0.74753 | 0.83559 | 0.75784 | 0.82528
Te, Tellurium (ppm) | 0.20005 | 0.16085 | 0.23926 | 0.16479 | 0.23531
Th, Thorium (ppm) | 14.8275 | 14.0493 | 15.6056 | 14.2357 | 15.4192
Ti, Titanium (wt.%) | 0.1106 | 0.10702 | 0.11418 | 0.10756 | 0.11364
Tl, Thallium (ppm) | 0.7211 | 0.69206 | 0.75013 | 0.69658 | 0.74562
Tm, Thulium (ppm) | 0.10679 | 0.087481 | 0.12611 | IND | IND
U, Uranium (ppm) | 5.0827 | 4.8618 | 5.3037 | 4.9505 | 5.215
V, Vanadium (ppm) | 3.8836 | 3.2788 | 4.4885 | 3.7535 | 4.0138
W, Tungsten (ppm) | 2.6107 | 2.4483 | 2.773 | 2.4593 | 2.762
Y, Yttrium (ppm) | 15.7894 | 15.138 | 16.4408 | 15.3491 | 16.2297
Yb, Ytterbium (ppm) | 0.6111 | 0.5258 | 0.6964 | 0.56523 | 0.65696
Zn, Zinc (ppm) | 146.8536 | 142.1406 | 151.5666 | 144.1815 | 149.5256
Zr, Zirconium (ppm) | 267.4545 | 258.942 | 275.9671 | 261.9472 | 272.9619
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 7.571 | 7.3313 | 7.8108 | 7.4573 | 7.6848
Ba, Barium (ppm) | 2841.4353 | 2717.5644 | 2965.3061 | 2802.418 | 2880.4525
Be, Beryllium (ppm) | 3.1472 | 2.503 | 3.7914 | IND | IND
Bi, Bismuth (ppm) | 10.4407 | 9.5594 | 11.322 | 10.085 | 10.7963
Ca, Calcium (wt.%) | 0.63822 | 0.6007 | 0.67575 | 0.61501 | 0.66144
Ce, Cerium (ppm) | 87.7798 | 79.7316 | 95.8279 | 84.8169 | 90.7427
Co, Cobalt (ppm) | 17.5452 | 15.8605 | 19.2298 | 16.9599 | 18.1304
Cs, Caesium (ppm) | 7.049 | 6.5316 | 7.5664 | 6.826 | 7.272
Cu, Copper (wt.%) | 0.31092 | 0.29958 | 0.32226 | 0.30449 | 0.31735
Dy, Dysprosium (ppm) | 3.7447 | 3.2374 | 4.252 | 3.5669 | 3.9224
Er, Erbium (ppm) | 1.1261 | 0.89896 | 1.3532 | 1.0528 | 1.1993
Eu, Europium (ppm) | 1.5121 | 1.3887 | 1.6355 | 1.4475 | 1.5767
Fe, Iron (wt.%) | 4.9154 | 4.7159 | 5.1149 | 4.8466 | 4.9842
Ga, Gallium (ppm) | 28.044 | 26.039 | 30.049 | 26.7735 | 29.3145
Gd, Gadolinium (ppm) | 6.0505 | 5.43 | 6.6709 | 5.7107 | 6.3902
Ge, Germanium (ppm) | 1.7396 | 0.95997 | 2.5191 | IND | IND
Hf, Hafnium (ppm) | 6.9067 | 5.9105 | 7.9028 | IND | IND
Ho, Holmium (ppm) | 0.52095 | 0.42641 | 0.61549 | 0.47759 | 0.56432
In, Indium (ppm) | 1.1642 | 0.99594 | 1.3326 | IND | IND
K, Potassium (wt.%) | 2.9834 | 2.8533 | 3.1135 | 2.9147 | 3.0521
La, Lanthanum (ppm) | 44.2188 | 41.5677 | 46.8698 | 42.317 | 46.1205
Li, Lithium (ppm) | 20.8233 | 18.769 | 22.8777 | 19.0583 | 22.5884
Mg, Magnesium (wt.%) | 0.23285 | 0.22105 | 0.24465 | 0.2246 | 0.2411
Mn, Manganese (wt.%) | 0.032176 | 0.03072 | 0.033632 | 0.03149 | 0.032861
Mo, Molybdenum (ppm) | 4.9994 | 3.8815 | 6.1173 | IND | IND
Nb, Niobium (ppm) | 17.6687 | 14.5861 | 20.7513 | 16.508 | 18.8294
Nd, Neodymium (ppm) | 38.3598 | 34.0943 | 42.6253 | 36.2876 | 40.432
P, Phosphorus (wt.%) | 0.028886 | 0.025018 | 0.032753 | IND | IND
Pb, Lead (ppm) | 31.3709 | 27.1936 | 35.5482 | 29.0867 | 33.6551
Pr, Praseodymium (ppm) | 10.1537 | 8.4961 | 11.8112 | 9.7881 | 10.5193
Rb, Rubidium (ppm) | 137.1498 | 132.8042 | 141.4954 | 134.3291 | 139.9704
Sb, Antimony (ppm) | 2.045 | 1.6971 | 2.393 | 1.828 | 2.262
Si, Silicon (wt.%) | 31.177 | 30.2877 | 32.0663 | 30.7528 | 31.6011
Sm, Samarium (ppm) | 7.1319 | 6.3301 | 7.9337 | 6.6209 | 7.6429
Sr, Strontium (ppm) | 163.6181 | 153.509 | 173.7272 | 159.7818 | 167.4545
Ta, Tantalum (ppm) | 1.5537 | 1.0648 | 2.0427 | 1.3997 | 1.7077
Tb, Terbium (ppm) | 0.78633 | 0.67003 | 0.90264 | 0.72033 | 0.85233
Th, Thorium (ppm) | 14.7302 | 13.7018 | 15.7586 | 14.2741 | 15.1863
Ti, Titanium (wt.%) | 0.11488 | 0.10634 | 0.12342 | 0.104 | 0.12576
Tl, Thallium (ppm) | 0.73883 | 0.62167 | 0.85598 | IND | IND
Tm, Thulium (ppm) | 0.1225 | 0.09772 | 0.14728 | IND | IND
U, Uranium (ppm) | 5.2216 | 4.9931 | 5.4502 | 5.0139 | 5.4294
Y, Yttrium (ppm) | 16.3803 | 13.7707 | 18.9899 | 15.7307 | 17.03
Yb, Ytterbium (ppm) | 0.67944 | 0.53567 | 0.82322 | IND | IND
Zn, Zinc (ppm) | 153.2494 | 131.8215 | 174.6773 | 145.0884 | 161.4103
Aqua Regia Digestion
Ag, Silver (ppm) | 0.67162 | 0.63284 | 0.71039 | 0.64056 | 0.70267
Al, Aluminium (wt.%) | 0.57522 | 0.54541 | 0.60504 | 0.55873 | 0.59171
As, Arsenic (ppm) | 11.8404 | 11.1543 | 12.5264 | 11.3722 | 12.3086
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 184.8151 | 172.6349 | 196.9954 | 178.8128 | 190.8175
Be, Beryllium (ppm) | 0.84692 | 0.80234 | 0.89149 | 0.81629 | 0.87755
Bi, Bismuth (ppm) | 9.9095 | 9.585 | 10.2339 | 9.6502 | 10.1687
Ca, Calcium (wt.%) | 0.36288 | 0.35081 | 0.37494 | 0.35564 | 0.37011
Cd, Cadmium (ppm) | 0.40096 | 0.36881 | 0.43311 | 0.37206 | 0.42986
Ce, Cerium (ppm) | 76.5835 | 73.249 | 79.9179 | 75.0393 | 78.1277
Co, Cobalt (ppm) | 15.8359 | 15.1666 | 16.5052 | 15.4498 | 16.222
Cr, Chromium (ppm) | 10.029 | 9.0306 | 11.0274 | 9.3102 | 10.7478
Cs, Caesium (ppm) | 0.97597 | 0.88424 | 1.0677 | 0.94582 | 1.0061
Cu, Copper (wt.%) | 0.31608 | 0.30743 | 0.32474 | 0.31208 | 0.32009
Fe, Iron (wt.%) | 4.2176 | 4.0882 | 4.3469 | 4.1483 | 4.2868
Ga, Gallium (ppm) | 6.9987 | 6.6377 | 7.3598 | 6.7696 | 7.2278
Ge, Germanium (ppm) | 0.10833 | 0.074082 | 0.14258 | IND | IND
Hf, Hafnium (ppm) | 1.136 | 1.03 | 1.242 | 1.0871 | 1.185
In, Indium (ppm) | 1.0703 | 1.0323 | 1.1083 | 1.0355 | 1.1051
K, Potassium (wt.%) | 0.22818 | 0.21449 | 0.24187 | 0.22011 | 0.23625
La, Lanthanum (ppm) | 37.5884 | 35.7753 | 39.4015 | 36.6876 | 38.4893
Li, Lithium (ppm) | 3.267 | 2.9545 | 3.5795 | 3.0622 | 3.4719
Mg, Magnesium (wt.%) | 0.097014 | 0.090147 | 0.10388 | 0.094609 | 0.099419
Mn, Manganese (wt.%) | 0.028375 | 0.027505 | 0.029245 | 0.027869 | 0.028881
Mo, Molybdenum (ppm) | 4.4654 | 4.2272 | 4.7036 | 4.2875 | 4.6433
Na, Sodium (wt.%) | 0.069706 | 0.061952 | 0.07746 | 0.066982 | 0.07243
Nb, Niobium (ppm) | 0.46902 | 0.35579 | 0.58225 | 0.43521 | 0.50282
Ni, Nickel (ppm) | 2.6921 | 2.3885 | 2.9958 | 2.5018 | 2.8824
P, Phosphorus (wt.%) | 0.024514 | 0.023462 | 0.025565 | 0.02364 | 0.025387
Pb, Lead (ppm) | 16.8719 | 16.0365 | 17.7073 | 16.4225 | 17.3212
Rb, Rubidium (ppm) | 12.5692 | 11.5688 | 13.5695 | 11.9593 | 13.179
S, Sulphur (wt.%) | 0.040411 | 0.038408 | 0.042413 | 0.039111 | 0.04171
Sb, Antimony (ppm) | 1.263 | 1.1786 | 1.3473 | 1.2222 | 1.3038
Sc, Scandium (ppm) | 1.3916 | 1.2652 | 1.518 | IND | IND
Se, Selenium (ppm) | 4.5199 | 4.0418 | 4.9981 | 4.2419 | 4.7979
Sn, Tin (ppm) | 1.3193 | 1.2315 | 1.4071 | 1.2728 | 1.3658
Sr, Strontium (ppm) | 10.7677 | 10.3407 | 11.1948 | 10.4004 | 11.1351
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.43206 | 0.3715 | 0.49262 | 0.40977 | 0.45435
Te, Tellurium (ppm) | 0.17518 | 0.13615 | 0.2142 | IND | IND
Th, Thorium (ppm) | 8.3231 | 7.8301 | 8.8161 | 8.0971 | 8.5491
Ti, Titanium (wt.%) | 0.015324 | 0.013675 | 0.016974 | 0.014275 | 0.016373
Tl, Thallium (ppm) | 0.072655 | 0.062957 | 0.082354 | IND | IND
U, Uranium (ppm) | 2.2312 | 2.0842 | 2.3781 | 2.1522 | 2.3101
V, Vanadium (ppm) | 2.0127 | 1.8987 | 2.1266 | IND | IND
W, Tungsten (ppm) | 0.66228 | 0.55991 | 0.76465 | 0.59916 | 0.7254
Y, Yttrium (ppm) | 6.6993 | 6.2191 | 7.1794 | 6.4924 | 6.9062
Yb, Ytterbium (ppm) | 0.23387 | 0.19817 | 0.26957 | IND | IND
Zn, Zinc (ppm) | 67.5471 | 65.0827 | 70.0116 | 65.3475 | 69.7467
Zr, Zirconium (ppm) | 46.3833 | 43.3337 | 49.4329 | 44.7884 | 47.9782
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 906b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 2.1667 | Pt | ppb | < 3
Infrared Combustion
C | wt.% | 0.12417 | S | wt.% | 0.06
Alkaline Leach
S-(Sulphate) | wt.% | 0.045 | S-(Sulphide) | wt.% | < 0.05
4-Acid Digestion
B | ppm | < 10 | Hg | ppm | < 1 | Pt | ppb | 11.3333
Peroxide Fusion ICP
Ag | ppm | 0.62083 | Lu | ppm | 0.068542 | Se | ppm | 5
As | ppm | 14.4028 | Na | wt.% | 0.8325 | Sn | ppm | 8.6608
B | ppm | 41.4167 | Ni | ppm | 8.0023 | Te | ppm | 2.1111
Cd | ppm | 0.33333 | Re | ppm | 0.0099167 | V | ppm | 5.6292
Cr | ppm | 34.3304 | S | wt.% | 0.046926 | W | ppm | 2.7944
Hg | ppm | < 0.1 | Sc | ppm | 6.33 | Zr | ppm | 278.4583
Aqua Regia Digestion
Dy | ppm | 1.7998 | Ho | ppm | 0.23098 | Pt | ppb | 3.3333
Er | ppm | 0.47267 | Lu | ppm | 0.031333 | Re | ppm | 0.0012593
Eu | ppm | 0.98551 | Nd | ppm | 32.0002 | Sm | ppm | 5.3405
Gd | ppm | 4.0711 | Pd | ppb | < 10 | Tm | ppm | 0.042833
Hg | ppm | 0.014806 | Pr | ppm | 8.6793
Borate Fusion XRF
Al2O3 | wt.% | 14.575 | Fe2O3 | wt.% | 7.2128 | S | wt.% | 0.045
As | ppm | 10 | K2O | wt.% | 3.525 | SiO2 | wt.% | 67
BaO | ppm | 3494.645 | MgO | wt.% | 0.405 | Sn | ppm | < 10
CaO | wt.% | 0.9 | MnO | wt.% | 0.04 | Sr | ppm | 170
Cl | ppm | 125 | Na2O | wt.% | 3.545 | TiO2 | wt.% | 0.2
Co | ppm | 10 | Ni | ppm | 20 | V2O5 | ppm | 26.778
Cr2O3 | ppm | 18.2688 | P2O5 | wt.% | 0.063019 | Zn | ppm | 140
Cu | wt.% | 0.3195 | Pb | ppm | 40 | Zr | ppm | 285
Thermogravimetry
LOI1000 | wt.% | 1.755
Laser Ablation ICP-MS
Ag | ppm | 0.7 | Hf | ppm | 7.705 | Sm | ppm | 7.685
As | ppm | 12.6 | Ho | ppm | 0.555 | Sn | ppm | 8.1
Ba | ppm | 2765 | In | ppm | 1.025 | Sr | ppm | 161
Be | ppm | 3.3 | La | ppm | 45.85 | Ta | ppm | 1.48
Bi | ppm | 10.25 | Lu | ppm | 0.085 | Tb | ppm | 0.845
Cd | ppm | 0.4 | Mn | wt.% | 0.0318 | Te | ppm | < 0.2
Ce | ppm | 91.05 | Mo | ppm | 5 | Th | ppm | 14.55
Co | ppm | 18.05 | Nb | ppm | 19 | Ti | wt.% | 0.1185
Cr | ppm | 11.5 | Nd | ppm | 38.5 | Tl | ppm | 0.4
Cs | ppm | 6.83 | Ni | ppm | 5 | Tm | ppm | 0.135
Cu | wt.% | 0.3245 | Pb | ppm | 31 | U | ppm | 5.13
Dy | ppm | 3.93 | Pr | ppm | 10.65 | V | ppm | 4.8
Er | ppm | 1.125 | Rb | ppm | 132.5 | W | ppm | 3.75
Eu | ppm | 1.47 | Re | ppm | < 0.01 | Y | ppm | 16.1
Ga | ppm | 27 | Sb | ppm | 2.1 | Yb | ppm | 0.705
Gd | ppm | 5.99 | Sc | ppm | 3.95 | Zn | ppm | 152.5
Ge | ppm | 1.325 | Se | ppm | < 5 | Zr | ppm | 283.5
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 906b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 52.3638 | 3.6452 | 45.0733 | 59.6542 | 41.4281 | 63.2995 | 6.961% | 13.92% | 20.88% | 49.7456 | 54.982
Aqua Regia Digestion (sample weights 10-50g)
Au, ppb | 51.8384 | 3.8282 | 44.1819 | 59.4948 | 40.3536 | 63.3231 | 7.385% | 14.77% | 22.15% | 49.2464 | 54.4303
Sequential Leach
Cu-Res(4A), wt.% | 0.038743 | 0.0023271 | 0.034088 | 0.043397 | 0.031761 | 0.045724 | 6.007% | 12.01% | 18.02% | 0.036806 | 0.04068
Cu-Sol(CN), wt.% | 0.0154 | 0.0029572 | 0.0094859 | 0.021315 | 0.0065288 | 0.024272 | 19.2% | 38.4% | 57.61% | 0.01463 | 0.01617
Cu-Sol(H2SO4), wt.% | 0.25975 | 0.0074626 | 0.24482 | 0.27467 | 0.23736 | 0.28214 | 2.873% | 5.746% | 8.619% | 0.24676 | 0.27274
Cu-Tot(Calc), wt.% | 0.31258 | 0.0071138 | 0.29836 | 0.32681 | 0.29124 | 0.33393 | 2.276% | 4.552% | 6.827% | 0.29696 | 0.32821
Gas / Liquid Pycnometry
SG, Unity | 2.7413 | 0.046914 | 2.6474 | 2.8351 | 2.6005 | 2.882 | 1.711% | 3.423% | 5.134% | 2.6042 | 2.8783
4-Acid Digestion
Ag, ppm | 0.70937 | 0.035829 | 0.63771 | 0.78103 | 0.60188 | 0.81686 | 5.051% | 10.1% | 15.15% | 0.6739 | 0.74484
Al, wt.% | 7.4554 | 0.26111 | 6.9332 | 7.9776 | 6.6721 | 8.2387 | 3.502% | 7.005% | 10.51% | 7.0826 | 7.8282
As, ppm | 13.6874 | 1.2702 | 11.1469 | 16.2279 | 9.8767 | 17.4981 | 9.28% | 18.56% | 27.84% | 13.003 | 14.3718
Ba, ppm | 2819.2182 | 85.1788 | 2648.8605 | 2989.5758 | 2563.6817 | 3074.7546 | 3.021% | 6.043% | 9.064% | 2678.2573 | 2960.1791
Be, ppm | 2.986 | 0.11728 | 2.7515 | 3.2206 | 2.6342 | 3.3378 | 3.928% | 7.855% | 11.78% | 2.8367 | 3.1353
Bi, ppm | 9.993 | 0.52546 | 8.9421 | 11.0439 | 8.4166 | 11.5694 | 5.258% | 10.52% | 15.77% | 9.4933 | 10.4926
Ca, wt.% | 0.64419 | 0.023546 | 0.59709 | 0.69128 | 0.57355 | 0.71483 | 3.655% | 7.31% | 10.97% | 0.61198 | 0.6764
Cd, ppm | 0.41753 | 0.039387 | 0.33876 | 0.49631 | 0.29937 | 0.53569 | 9.433% | 18.87% | 28.3% | 0.39665 | 0.43841
Ce, ppm | 92.1932 | 5.4835 | 81.2263 | 103.1601 | 75.7428 | 108.6435 | 5.948% | 11.9% | 17.84% | 87.5835 | 96.8028
Co, ppm | 17.5817 | 0.78578 | 16.0101 | 19.1532 | 15.2244 | 19.939 | 4.469% | 8.939% | 13.41% | 16.7026 | 18.4608
Cr, ppm | 10.1781 | 1.1384 | 7.9013 | 12.455 | 6.7629 | 13.5934 | 11.18% | 22.37% | 33.55% | 9.6692 | 10.6871
Cs, ppm | 6.97 | 0.20071 | 6.5686 | 7.3715 | 6.3679 | 7.5722 | 2.88% | 5.759% | 8.639% | 6.6215 | 7.3185
Cu, wt.% | 0.31434 | 0.0083745 | 0.29759 | 0.33109 | 0.28922 | 0.33947 | 2.664% | 5.328% | 7.992% | 0.29863 | 0.33006
Dy, ppm | 3.705 | 0.2475 | 3.21 | 4.2 | 2.9625 | 4.4475 | 6.68% | 13.36% | 20.04% | 3.5198 | 3.8903
Er, ppm | 1.0734 | 0.07734 | 0.91877 | 1.2281 | 0.84143 | 1.3055 | 7.205% | 14.41% | 21.61% | 1.0198 | 1.1271
Eu, ppm | 1.5052 | 0.19272 | 1.1198 | 1.8907 | 0.92709 | 2.0834 | 12.8% | 25.61% | 38.41% | 1.43 | 1.5805
Fe, wt.% | 4.8529 | 0.086456 | 4.68 | 5.0258 | 4.5935 | 5.1123 | 1.782% | 3.563% | 5.345% | 4.6103 | 5.0956
Ga, ppm | 28.1329 | 1.0015 | 26.1299 | 30.1359 | 25.1284 | 31.1374 | 3.56% | 7.12% | 10.68% | 26.7262 | 29.5395
Gd, ppm | 5.9026 | 0.57614 | 4.7503 | 7.0549 | 4.1742 | 7.631 | 9.761% | 19.52% | 29.28% | 5.6075 | 6.1977
Ge, ppm | 0.20724 | 0.05498 | 0.097277 | 0.3172 | 0.042297 | 0.37218 | 26.53% | 53.06% | 79.59% | 0.19688 | 0.2176
Hf, ppm | 7.1516 | 0.32944 | 6.4927 | 7.8105 | 6.1633 | 8.1399 | 4.607% | 9.213% | 13.82% | 6.794 | 7.5092
Ho, ppm | 0.50291 | 0.069756 | 0.3634 | 0.64242 | 0.29365 | 0.71218 | 13.87% | 27.74% | 41.61% | 0.47777 | 0.52806
In, ppm | 1.1595 | 0.075022 | 1.0095 | 1.3095 | 0.93444 | 1.3846 | 6.47% | 12.94% | 19.41% | 1.1015 | 1.2175
K, wt.% | 2.8848 | 0.067668 | 2.7494 | 3.0201 | 2.6817 | 3.0878 | 2.346% | 4.691% | 7.037% | 2.7405 | 3.029
La, ppm | 44.9691 | 2.9668 | 39.0354 | 50.9027 | 36.0686 | 53.8696 | 6.597% | 13.19% | 19.79% | 42.7206 | 47.2175
Li, ppm | 20.5939 | 0.9275 | 18.7389 | 22.4489 | 17.8114 | 23.3764 | 4.504% | 9.008% | 13.51% | 19.5642 | 21.6236
Lu, ppm | 0.084998 | 0.0049037 | 0.075191 | 0.094806 | 0.070287 | 0.099709 | 5.769% | 11.54% | 17.31% | 0.080748 | 0.089248
Mg, wt.% | 0.23304 | 0.012772 | 0.2075 | 0.25859 | 0.19473 | 0.27136 | 5.48% | 10.96% | 16.44% | 0.22139 | 0.2447
Mn, wt.% | 0.032119 | 0.0014743 | 0.02917 | 0.035067 | 0.027696 | 0.036541 | 4.59% | 9.18% | 13.77% | 0.030513 | 0.033724
Mo, ppm | 4.9125 | 0.36646 | 4.1796 | 5.6454 | 3.8131 | 6.0119 | 7.46% | 14.92% | 22.38% | 4.6669 | 5.1581
Na, wt.% | 2.6259 | 0.13321 | 2.3595 | 2.8923 | 2.2263 | 3.0255 | 5.073% | 10.15% | 15.22% | 2.4946 | 2.7572
Nb, ppm | 18.2642 | 1.0773 | 16.1097 | 20.4187 | 15.0324 | 21.496 | 5.898% | 11.8% | 17.69% | 17.351 | 19.1774
Nd, ppm | 38.6178 | 3.4993 | 31.6192 | 45.6163 | 28.1199 | 49.1156 | 9.061% | 18.12% | 27.18% | 36.6869 | 40.5487
Ni, ppm | 3.2247 | 0.44321 | 2.3383 | 4.1111 | 1.8951 | 4.5544 | 13.74% | 27.49% | 41.23% | 3.0635 | 3.386
P, wt.% | 0.028074 | 0.0010554 | 0.025963 | 0.030185 | 0.024908 | 0.03124 | 3.759% | 7.519% | 11.28% | 0.02667 | 0.029478
Pb, ppm | 31.6988 | 1.3736 | 28.9516 | 34.4461 | 27.578 | 35.8197 | 4.333% | 8.667% | 13% | 30.1139 | 33.2838
Pr, ppm | 10.1972 | 0.61935 | 8.9585 | 11.4359 | 8.3391 | 12.0552 | 6.074% | 12.15% | 18.22% | 9.6873 | 10.707
Rb, ppm | 138.48 | 6.8927 | 124.6945 | 152.2654 | 117.8017 | 159.1582 | 4.977% | 9.955% | 14.93% | 131.556 | 145.404
Re, ppm | 0.0022146 | 0.00061862 | 0.00097735 | 0.0034518 | 0.00035873 | 0.0040704 | 27.93% | 55.87% | 83.8% | 0.0021039 | 0.0023253
S, wt.% | 0.044263 | 0.0057232 | 0.032817 | 0.055709 | 0.027093 | 0.061432 | 12.93% | 25.86% | 38.79% | 0.04205 | 0.046476
Sb, ppm | 2.1562 | 0.14342 | 1.8693 | 2.443 | 1.7259 | 2.5864 | 6.651% | 13.3% | 19.95% | 2.0484 | 2.264
Sc, ppm | 4.7135 | 0.54276 | 3.628 | 5.799 | 3.0852 | 6.3418 | 11.52% | 23.03% | 34.55% | 4.4778 | 4.9491
Se, ppm | 5.1088 | 0.50815 | 4.0925 | 6.1251 | 3.5844 | 6.6333 | 9.947% | 19.89% | 29.84% | 4.8534 | 5.3642
Sm, ppm | 7.3708 | 0.52487 | 6.3211 | 8.4205 | 5.7962 | 8.9454 | 7.121% | 14.24% | 21.36% | 7.0023 | 7.7393
Sn, ppm | 4.2394 | 0.2498 | 3.7398 | 4.739 | 3.4901 | 4.9888 | 5.892% | 11.78% | 17.68% | 4.0275 | 4.4514
Sr, ppm | 164.7968 | 6.2779 | 152.241 | 177.3526 | 145.9631 | 183.6305 | 3.809% | 7.619% | 11.43% | 156.557 | 173.0366
Ta, ppm | 1.3654 | 0.13292 | 1.0996 | 1.6312 | 0.96664 | 1.7641 | 9.735% | 19.47% | 29.2% | 1.2971 | 1.4337
Tb, ppm | 0.79156 | 0.023088 | 0.74538 | 0.83774 | 0.7223 | 0.86082 | 2.917% | 5.834% | 8.75% | 0.75198 | 0.83114
Te, ppm | 0.20005 | 0.028977 | 0.1421 | 0.25801 | 0.11312 | 0.28698 | 14.48% | 28.97% | 43.45% | 0.19005 | 0.21006
Th, ppm | 14.8275 | 0.68902 | 13.4494 | 16.2055 | 12.7604 | 16.8945 | 4.647% | 9.294% | 13.94% | 14.0861 | 15.5688
Ti, wt.% | 0.1106 | 0.0048621 | 0.10088 | 0.12033 | 0.096015 | 0.12519 | 4.396% | 8.792% | 13.19% | 0.10507 | 0.11613
Tl, ppm | 0.7211 | 0.031652 | 0.6578 | 0.7844 | 0.62614 | 0.81605 | 4.389% | 8.779% | 13.17% | 0.68504 | 0.75715
Tm, ppm | 0.10679 | 0.017879 | 0.071036 | 0.14255 | 0.053157 | 0.16043 | 16.74% | 33.48% | 50.22% | 0.10145 | 0.11213
U, ppm | 5.0827 | 0.20859 | 4.6656 | 5.4999 | 4.457 | 5.7085 | 4.104% | 8.208% | 12.31% | 4.8286 | 5.3369
V, ppm | 3.8836 | 0.54667 | 2.7903 | 4.977 | 2.2436 | 5.5236 | 14.08% | 28.15% | 42.23% | 3.6895 | 4.0778
W, ppm | 2.6107 | 0.15918 | 2.2923 | 2.929 | 2.1331 | 3.0882 | 6.097% | 12.19% | 18.29% | 2.4801 | 2.7412
Y, ppm | 15.7894 | 0.64437 | 14.5007 | 17.0781 | 13.8563 | 17.7225 | 4.081% | 8.162% | 12.24% | 14.9999 | 16.5789
Yb, ppm | 0.6111 | 0.065149 | 0.4808 | 0.7414 | 0.41565 | 0.80655 | 10.66% | 21.32% | 31.98% | 0.58054 | 0.64165
Zn, ppm | 146.8536 | 5.325 | 136.2035 | 157.5036 | 130.8785 | 162.8287 | 3.626% | 7.252% | 10.88% | 139.5109 | 154.1962
Zr, ppm | 267.4545 | 10.3741 | 246.7063 | 288.2028 | 236.3322 | 298.5769 | 3.879% | 7.758% | 11.64% | 254.0818 | 280.8273
Peroxide Fusion ICP
Al, wt.% | 7.571 | 0.23255 | 7.1059 | 8.0361 | 6.8734 | 8.2687 | 3.072% | 6.143% | 9.215% | 7.1925 | 7.9496
Ba, ppm | 2841.4353 | 117.6179 | 2606.1994 | 3076.6711 | 2488.5815 | 3194.289 | 4.139% | 8.279% | 12.42% | 2699.3635 | 2983.507
Be, ppm | 3.1472 | 0.39017 | 2.3668 | 3.9275 | 1.9767 | 4.3177 | 12.4% | 24.79% | 37.19% | 2.9898 | 3.3045
Bi, ppm | 10.4407 | 0.38533 | 9.67 | 11.2113 | 9.2847 | 11.5967 | 3.691% | 7.381% | 11.07% | 9.9186 | 10.9627
Ca, wt.% | 0.63822 | 0.040744 | 0.55674 | 0.71971 | 0.51599 | 0.76045 | 6.384% | 12.77% | 19.15% | 0.60631 | 0.67013
Ce, ppm | 87.7798 | 5.8931 | 75.9936 | 99.566 | 70.1005 | 105.4591 | 6.714% | 13.43% | 20.14% | 83.3908 | 92.1688
Co, ppm | 17.5452 | 1.9128 | 13.7196 | 21.3707 | 11.8069 | 23.2834 | 10.9% | 21.8% | 32.71% | 16.6679 | 18.4224
Cs, ppm | 7.049 | 0.4217 | 6.2056 | 7.8924 | 5.7839 | 8.3141 | 5.982% | 11.96% | 17.95% | 6.6966 | 7.4015
Cu, wt.% | 0.31092 | 0.016763 | 0.2774 | 0.34445 | 0.26063 | 0.36121 | 5.391% | 10.78% | 16.17% | 0.29538 | 0.32647
Dy, ppm | 3.7447 | 0.19951 | 3.3457 | 4.1437 | 3.1461 | 4.3432 | 5.328% | 10.66% | 15.98% | 3.5574 | 3.9319
Er, ppm | 1.1261 | 0.097478 | 0.9311 | 1.321 | 0.83362 | 1.4185 | 8.657% | 17.31% | 25.97% | 1.0698 | 1.1824
Eu, ppm | 1.5121 | 0.078731 | 1.3546 | 1.6696 | 1.2759 | 1.7483 | 5.207% | 10.41% | 15.62% | 1.4365 | 1.5877
Fe, wt.% | 4.9154 | 0.23203 | 4.4513 | 5.3795 | 4.2193 | 5.6115 | 4.721% | 9.441% | 14.16% | 4.6696 | 5.1612
Ga, ppm | 28.044 | 1.9259 | 24.1921 | 31.8958 | 22.2662 | 33.8217 | 6.867% | 13.73% | 20.6% | 26.6418 | 29.4462
Gd, ppm | 6.0505 | 0.58161 | 4.8873 | 7.2137 | 4.3057 | 7.7953 | 9.613% | 19.23% | 28.84% | 5.748 | 6.353
Ge, ppm | 1.7396 | 0.33934 | 1.0609 | 2.4182 | 0.72154 | 2.7576 | 19.51% | 39.01% | 58.52% | 1.6526 | 1.8265
Hf, ppm | 6.9067 | 0.4849 | 5.9369 | 7.8765 | 5.452 | 8.3614 | 7.021% | 14.04% | 21.06% | 6.5613 | 7.252
Ho, ppm | 0.52095 | 0.045308 | 0.43034 | 0.61157 | 0.38503 | 0.65688 | 8.697% | 17.39% | 26.09% | 0.4949 | 0.547
In, ppm | 1.1642 | 0.092088 | 0.98007 | 1.3484 | 0.88798 | 1.4405 | 7.91% | 15.82% | 23.73% | 1.106 | 1.2225
K, wt.% | 2.9834 | 0.19522 | 2.593 | 3.3739 | 2.3977 | 3.5691 | 6.544% | 13.09% | 19.63% | 2.8342 | 3.1326
La, ppm | 44.2188 | 1.99 | 40.2387 | 48.1988 | 38.2487 | 50.1888 | 4.5% | 9.001% | 13.5% | 42.0078 | 46.4297
Li, ppm | 20.8233 | 2.8113 | 15.2008 | 26.4459 | 12.3895 | 29.2571 | 13.5% | 27% | 40.5% | 19.7822 | 21.8645
Mg, wt.% | 0.23285 | 0.011692 | 0.20946 | 0.25623 | 0.19777 | 0.26793 | 5.021% | 10.04% | 15.06% | 0.22121 | 0.24449
Mn, wt.% | 0.032176 | 0.0020427 | 0.02809 | 0.036261 | 0.026048 | 0.038304 | 6.348% | 12.7% | 19.05% | 0.030567 | 0.033785
Mo, ppm | 4.9994 | 0.82431 | 3.3508 | 6.6481 | 2.5265 | 7.4724 | 16.49% | 32.98% | 49.46% | 4.7495 | 5.2494
Nb, ppm | 17.6687 | 2.75 | 12.1688 | 23.1686 | 9.4188 | 25.9186 | 15.56% | 31.13% | 46.69% | 16.7853 | 18.5521
Nd, ppm | 38.3598 | 3.7321 | 30.8957 | 45.8239 | 27.1636 | 49.556 | 9.729% | 19.46% | 29.19% | 36.4418 | 40.2778
P, wt.% | 0.028886 | 0.0070126 | 0.01486 | 0.042911 | 0.0078479 | 0.049923 | 24.28% | 48.55% | 72.83% | 0.027441 | 0.03033
Pb, ppm | 31.3709 | 4.0465 | 23.2778 | 39.464 | 19.2313 | 43.5105 | 12.9% | 25.8% | 38.7% | 29.8024 | 32.9394
Pr, ppm | 10.1537 | 0.63138 | 8.8909 | 11.4164 | 8.2595 | 12.0478 | 6.218% | 12.44% | 18.65% | 9.646 | 10.6614
Rb, ppm | 137.1498 | 3.5337 | 130.0824 | 144.2172 | 126.5487 | 147.7509 | 2.577% | 5.153% | 7.73% | 130.2923 | 144.0073
Sb, ppm | 2.045 | 0.2649 | 1.5152 | 2.5748 | 1.2503 | 2.8397 | 12.95% | 25.91% | 38.86% | 1.9428 | 2.1473
Si, wt.% | 31.177 | 0.88718 | 29.4026 | 32.9513 | 28.5154 | 33.8385 | 2.846% | 5.691% | 8.537% | 29.6181 | 32.7358
Sm, ppm | 7.1319 | 0.70295 | 5.726 | 8.5378 | 5.023 | 9.2408 | 9.856% | 19.71% | 29.57% | 6.7753 | 7.4885
Sr, ppm | 163.6181 | 11.5996 | 140.419 | 186.8173 | 128.8194 | 198.4168 | 7.089% | 14.18% | 21.27% | 155.4372 | 171.799
Ta, ppm | 1.5537 | 0.3904 | 0.77295 | 2.3345 | 0.38256 | 2.7249 | 25.13% | 50.25% | 75.38% | 1.4761 | 1.6314
Tb, ppm | 0.78633 | 0.073319 | 0.63969 | 0.93297 | 0.56637 | 1.0063 | 9.324% | 18.65% | 27.97% | 0.74702 | 0.82565
Th, ppm | 14.7302 | 0.48749 | 13.7552 | 15.7052 | 13.2677 | 16.1927 | 3.309% | 6.619% | 9.928% | 13.9937 | 15.4667
Ti, wt.% | 0.11488 | 0.0060126 | 0.10285 | 0.1269 | 0.096841 | 0.13292 | 5.234% | 10.47% | 15.7% | 0.10914 | 0.12062
Tl, ppm | 0.73883 | 0.061762 | 0.6153 | 0.86235 | 0.55354 | 0.92411 | 8.36% | 16.72% | 25.08% | 0.70188 | 0.77577
Tm, ppm | 0.1225 | 0.027607 | 0.067286 | 0.17771 | 0.039679 | 0.20532 | 22.54% | 45.07% | 67.61% | 0.11637 | 0.12862
U, ppm | 5.2216 | 0.2315 | 4.7586 | 5.6846 | 4.5271 | 5.9161 | 4.434% | 8.867% | 13.3% | 4.9605 | 5.4827
Y, ppm | 16.3803 | 1.3659 | 13.6484 | 19.1122 | 12.2825 | 20.4782 | 8.339% | 16.68% | 25.02% | 15.5613 | 17.1993
Yb, ppm | 0.67944 | 0.060424 | 0.5586 | 0.80029 | 0.49817 | 0.86072 | 8.893% | 17.79% | 26.68% | 0.64547 | 0.71342
Zn, ppm | 153.2494 | 17.5847 | 118.0799 | 188.4188 | 100.4952 | 206.0035 | 11.47% | 22.95% | 34.42% | 145.5869 | 160.9118
Aqua Regia Digestion
Ag, ppm | 0.67162 | 0.04217 | 0.58728 | 0.75595 | 0.54511 | 0.79812 | 6.279% | 12.56% | 18.84% | 0.63803 | 0.7052
Al, wt.% | 0.57522 | 0.040026 | 0.49517 | 0.65527 | 0.45514 | 0.6953 | 6.958% | 13.92% | 20.88% | 0.54646 | 0.60398
As, ppm | 11.8404 | 0.67534 | 10.4897 | 13.1911 | 9.8144 | 13.8664 | 5.704% | 11.41% | 17.11% | 11.2484 | 12.4324
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 184.8151 | 21.1558 | 142.5036 | 227.1266 | 121.3479 | 248.2824 | 11.45% | 22.89% | 34.34% | 175.5744 | 194.0559
Be, ppm | 0.84692 | 0.042294 | 0.76233 | 0.93151 | 0.72004 | 0.9738 | 4.994% | 9.988% | 14.98% | 0.80457 | 0.88926
Bi, ppm | 9.9095 | 0.54458 | 8.8203 | 10.9986 | 8.2757 | 11.5432 | 5.496% | 10.99% | 16.49% | 9.414 | 10.4049
Ca, wt.% | 0.36288 | 0.016123 | 0.33063 | 0.39512 | 0.31451 | 0.41124 | 4.443% | 8.886% | 13.33% | 0.34473 | 0.38102
Cd, ppm | 0.40096 | 0.022893 | 0.35517 | 0.44674 | 0.33228 | 0.46964 | 5.709% | 11.42% | 17.13% | 0.38091 | 0.42101
Ce, ppm | 76.5835 | 5.4355 | 65.7124 | 87.4546 | 60.2768 | 92.8901 | 7.098% | 14.2% | 21.29% | 72.7543 | 80.4127
Co, ppm | 15.8359 | 0.61293 | 14.6101 | 17.0618 | 13.9971 | 17.6747 | 3.87% | 7.741% | 11.61% | 15.0441 | 16.6277
Cr, ppm | 10.029 | 0.75748 | 8.5141 | 11.544 | 7.7566 | 12.3015 | 7.553% | 15.11% | 22.66% | 9.5276 | 10.5305
Cs, ppm | 0.97597 | 0.14629 | 0.68339 | 1.2686 | 0.5371 | 1.4148 | 14.99% | 29.98% | 44.97% | 0.92717 | 1.0248
Cu, wt.% | 0.31608 | 0.010573 | 0.29494 | 0.33723 | 0.28437 | 0.3478 | 3.345% | 6.69% | 10.03% | 0.30028 | 0.33189
Fe, wt.% | 4.2176 | 0.13576 | 3.946 | 4.4891 | 3.8103 | 4.6248 | 3.219% | 6.438% | 9.657% | 4.0067 | 4.4284
Ga, ppm | 6.9987 | 0.35841 | 6.2819 | 7.7155 | 5.9235 | 8.074 | 5.121% | 10.24% | 15.36% | 6.6488 | 7.3487
Ge, ppm | 0.10833 | 0.026458 | 0.055418 | 0.16125 | 0.028961 | 0.18771 | 24.42% | 48.84% | 73.27% | 0.10292 | 0.11375
Hf, ppm | 1.136 | 0.122 | 0.892 | 1.38 | 0.77 | 1.502 | 10.74% | 21.48% | 32.22% | 1.0792 | 1.1928
In, ppm | 1.0703 | 0.036479 | 0.99734 | 1.1433 | 0.96086 | 1.1797 | 3.408% | 6.817% | 10.22% | 1.0168 | 1.1238
K, wt.% | 0.22818 | 0.022073 | 0.18403 | 0.27232 | 0.16196 | 0.2944 | 9.674% | 19.35% | 29.02% | 0.21677 | 0.23959
La, ppm | 37.5884 | 3.204 | 31.1804 | 43.9964 | 27.9764 | 47.2004 | 8.524% | 17.05% | 25.57% | 35.709 | 39.4679
Li, ppm | 3.267 | 0.3278 | 2.6114 | 3.9226 | 2.2836 | 4.2504 | 10.03% | 20.07% | 30.1% | 3.1037 | 3.4304
Mg, wt.% | 0.097014 | 0.0068958 | 0.083222 | 0.11081 | 0.076327 | 0.1177 | 7.108% | 14.22% | 21.32% | 0.092163 | 0.10186
Mn, wt.% | 0.028375 | 0.0011701 | 0.026035 | 0.030715 | 0.024865 | 0.031885 | 4.124% | 8.248% | 12.37% | 0.026956 | 0.029794
Mo, ppm | 4.4654 | 0.33743 | 3.7905 | 5.1402 | 3.4531 | 5.4777 | 7.556% | 15.11% | 22.67% | 4.2421 | 4.6887
Na, wt.% | 0.069706 | 0.0070419 | 0.055622 | 0.08379 | 0.048581 | 0.090832 | 10.1% | 20.2% | 30.31% | 0.066221 | 0.073192
Nb, ppm | 0.46902 | 0.17331 | 0.1224 | 0.81563 | 0 | 0.98894 | 36.95% | 73.9% | 110.9% | 0.44557 | 0.49247
Ni, ppm | 2.6921 | 0.19719 | 2.2977 | 3.0865 | 2.1005 | 3.2837 | 7.325% | 14.65% | 21.97% | 2.5575 | 2.8267
P, wt.% | 0.024514 | 0.0012062 | 0.022101 | 0.026926 | 0.020895 | 0.028132 | 4.921% | 9.841% | 14.76% | 0.023288 | 0.025739
Pb, ppm | 16.8719 | 0.89333 | 15.0852 | 18.6585 | 14.1919 | 19.5519 | 5.295% | 10.59% | 15.88% | 16.0283 | 17.7155
Rb, ppm | 12.5692 | 1.3557 | 9.8577 | 15.2806 | 8.502 | 16.6363 | 10.79% | 21.57% | 32.36% | 11.9407 | 13.1976
S, wt.% | 0.040411 | 0.0017252 | 0.03696 | 0.043861 | 0.035235 | 0.045586 | 4.269% | 8.539% | 12.81% | 0.03839 | 0.042431
Sb, ppm | 1.263 | 0.13315 | 0.99667 | 1.5293 | 0.86352 | 1.6624 | 10.54% | 21.09% | 31.63% | 1.1998 | 1.3261
Sc, ppm | 1.3916 | 0.18258 | 1.0264 | 1.7568 | 0.84387 | 1.9393 | 13.12% | 26.24% | 39.36% | 1.322 | 1.4612
Se, ppm | 4.5199 | 0.36908 | 3.7818 | 5.2581 | 3.4127 | 5.6272 | 8.166% | 16.33% | 24.5% | 4.2939 | 4.7459
Sn, ppm | 1.3193 | 0.063859 | 1.1916 | 1.447 | 1.1277 | 1.5109 | 4.84% | 9.681% | 14.52% | 1.2534 | 1.3853
Sr, ppm | 10.7677 | 0.49248 | 9.7828 | 11.7527 | 9.2903 | 12.2452 | 4.574% | 9.147% | 13.72% | 10.2293 | 11.3061
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.43206 | 0.040751 | 0.35056 | 0.51357 | 0.30981 | 0.55432 | 9.432% | 18.86% | 28.3% | 0.41046 | 0.45367
Te, ppm | 0.17518 | 0.021912 | 0.13135 | 0.219 | 0.10944 | 0.24091 | 12.51% | 25.02% | 37.52% | 0.16642 | 0.18394
Th, ppm | 8.3231 | 0.67779 | 6.9675 | 9.6787 | 6.2897 | 10.3565 | 8.143% | 16.29% | 24.43% | 7.907 | 8.7393
Ti, wt.% | 0.015324 | 0.0031011 | 0.0091221 | 0.021526 | 0.0060209 | 0.024628 | 20.24% | 40.47% | 60.71% | 0.014558 | 0.01609
Tl, ppm | 0.072655 | 0.0091473 | 0.054361 | 0.09095 | 0.045213 | 0.1001 | 12.59% | 25.18% | 37.77% | 0.069022 | 0.076288
U, ppm | 2.2312 | 0.21532 | 1.8005 | 2.6618 | 1.5852 | 2.8771 | 9.651% | 19.3% | 28.95% | 2.1196 | 2.3427
V, ppm | 2.0127 | 0.072205 | 1.8682 | 2.1571 | 1.796 | 2.2293 | 3.588% | 7.175% | 10.76% | 1.912 | 2.1133
W, ppm | 0.66228 | 0.12417 | 0.41394 | 0.91062 | 0.28978 | 1.0348 | 18.75% | 37.5% | 56.25% | 0.62917 | 0.69539
Y, ppm | 6.6993 | 0.81081 | 5.0777 | 8.3209 | 4.2669 | 9.1317 | 12.1% | 24.21% | 36.31% | 6.3643 | 7.0343
Yb, ppm | 0.23387 | 0.042442 | 0.14899 | 0.31875 | 0.10654 | 0.3612 | 18.15% | 36.3% | 54.44% | 0.22218 | 0.24556
Zn, ppm | 67.5471 | 3.3574 | 60.8324 | 74.2619 | 57.475 | 77.6193 | 4.97% | 9.941% | 14.91% | 64.1698 | 70.9245
Zr, ppm | 46.3833 | 4.3262 | 37.7308 | 55.0357 | 33.4046 | 59.3619 | 9.327% | 18.65% | 27.98% | 44.0641 | 48.7024
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 906b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AL-5S*AAS | 5% sulphuric acid laech with atomic absorption spectroscopy
AL-5S*OES | 5% sulphuric acid leach with inductively coupled plasma optical emission spectroscopy
AL-5S*OES/AAS | 5% sulphuric acid leach with AAS or ICP-OES finish
AL-NaCN*AAS | sodium cyanide leach with atomic absorption spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
GASPYC | gas expansion pycnometry
INAA | instrumental neutron activation analysis
IRC | infrared combustion
IRC-x*Calc | calculation by difference from IRC total value
LIQPYC | liquid pycnometry
LOI*TGA | loss on ignition by thermogravimetric analysis
Na2CO3*GRAV | sodium carbonate solution with Gravimetric finish
Na2CO3*IRC | sodium carbonate leach with infrared combustion
NaCN*AAS | sodium cyanide with AAS finish
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
SQL-Tot*AAS | sequential leach analysis for total value method calculation with AAS finish
SQL-Tot*Calc | total value method calculated from sequential leach analysis
Text Values:
- | Not Analysed (Lab 2.04)
-- | Element Not Determined (Lab 2.14)
N.A. | Not Applicable (Lab 2.24)
N.A. | Not Applicable (Lab 2.27)
NA | Not Applicable (Lab 2.02)
NA | Not Applicable (Lab 2.23)
NA | Unable to report due to QC failure (Lab 2.26)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 906b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Labwest Minerals Analysis, Perth, WA, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS, Randfontein, Gauteng, South Africa
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 906b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppb) | Upscaled Value (ppb)
1 | 1 | ANSLu | 1 | 9 | 9 | 270669 | 0.085093 | 51.4656 | 53.1567
1 | 1 | ANSLu | 1 | 2 | 1 | 270670 | 0.087221 | 56.4188 | 53.4203
1 | 1 | ANSLu | 1 | 16 | 6 | 270671 | 0.08378 | 63.8765 | 53.8172
1 | 1 | ANSLu | 1 | 4 | 3 | 270672 | 0.083136 | 76.069 | 54.4662
1 | 1 | ANSLu | 1 | 8 | 7 | 270673 | 0.083146 | 43.0769 | 52.7102
1 | 1 | ANSLu | 1 | 10 | 9 | 270674 | 0.084272 | 47.6786 | 52.9551
1 | 1 | ANSLu | 1 | 17 | 8 | 270675 | 0.08736 | 43.9479 | 52.7565
1 | 1 | ANSLu | 1 | 11 | 2 | 270676 | 0.084792 | 69.8264 | 54.1339
1 | 1 | ANSLu | 1 | 13 | 4 | 270677 | 0.087333 | 53.9076 | 53.2866
1 | 1 | ANSLu | 1 | 3 | 3 | 270678 | 0.084887 | 51.3899 | 53.1526
1 | 1 | ANSLu | 1 | 5 | 5 | 270679 | 0.083387 | 50.1097 | 53.0845
1 | 1 | ANSLu | 1 | 7 | 7 | 270680 | 0.08615 | 49.2208 | 53.0372
1 | 1 | ANSLu | 1 | 6 | 5 | 270681 | 0.085347 | 46.5944 | 52.8974
1 | 1 | ANSLu | 1 | 14 | 4 | 270682 | 0.084853 | 48.3534 | 52.991
1 | 1 | ANSLu | 1 | 15 | 6 | 270683 | 0.084823 | 64.8784 | 53.8706
1 | 1 | ANSLu | 1 | 1 | 1 | 270684 | 0.08444 | 67.506 | 54.0104
1 | 1 | ANSLu | 1 | 18 | 8 | 270685 | 0.086636 | 37.7142 | 52.4248
1 | 1 | ANSLu | 1 | 19 | 10 | 270686 | 0.08349 | 50.0308 | 53.0803
1 | 1 | ANSLu | 1 | 20 | 10 | 270687 | 0.084196 | 48.3581 | 52.9913
1 | 1 | ANSLu | 1 | 12 | 2 | 270688 | 0.085414 | 44.6119 | 52.7919
Mean | [6 empty] | 0.084988 | 53.2517 | 53.2517
Median | [7 empty] | 50.0703 | 53.0824
SD | [7 empty] | 10.0778 | 0.53639
RSD | [7 empty] | 18.92% | 1.007%
RSD Upscale Factor | [7 empty] | 0.053225
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 906b (Certified Value 52.4 ppb) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.23 | 2.24 | 2.26 | 2.27 | ppb
|  |  | INAA | FA*MS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | 1
|  |  | 0.085g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 50g | 30g | 30g | 30g | 2
| 1 | 1 | 67.506 | 50 | 62 | 54.3614 | 49 | 52 | 51 | 51 | 48 | 50 | 55 | 51 | 53 | 54 | 48 | 55.6123 | 54 | 58 | 36.8 | 50 | 49 | 50 | 50 | 54 | 1
| 1 | 2 | 56.4188 | 50 | 52 | 53.7231 | 50 | 55.3333 | 52 | 55 | 50 | 51 | 55 | 49 | 50 | 41 | 48 | 58.5721 | 52 | 61 | 52.44 | 50 | 46 | 52 | 51 | 53 | #N/A
| 1 | 3 | 51.3899 | 63 | 56 | 53.1117 | 51 | 48.6667 | 52 | 53 | 51 | 51 | 52 | 51 | 50 | 43 | 43 | 55.8216 | 61 | 53 | 43.24 | 60 | 48 | 54 | 49 | 56 | 16
| 1 | 4 | 76.069 | 57 | 62 | 56.364 | 51 | 57.3333 | 51 | 58 | 51 | 51 | 53 | 55 | 52 | 50 | 47 | 57.5866 | 53 | 58 | 38.64 | 50 | 51 | 50 | 49 | 53 | 52.3638
| 1 | 5 | 50.1097 | 51 | 58 | 54.9725 | 52 | 49.3333 | 52 | 59 | 50 | 50 | 55 | 53 | 53 | 46 | 48 | 56.7015 | 54 | 51 | 45.08 | 50 | 52 | 48 | 51 | 54 | 7
| 1 | 6 | 46.5944 | 53 | 54 | 56.2359 | 53 | 52.3333 | 51 | 50 | 54 | 53 | 54 | 52 | 53 | 44 | 46 | 55.3718 | 52 | 52 | 39.56 | 50 | 49 | 54 | 50 | 53
|  | 7 | 49.2208
|  | 8 | 43.0769
|  | 9 | 51.4656
|  | 10 | 47.6786
|  | 11 | 69.8264
|  | 12 | 44.6119
|  | 13 | 53.9076
|  | 14 | 48.3534
|  | 15 | 64.8784
|  | 16 | 63.8765
|  | 17 | 43.9479
|  | 18 | 37.7142
|  | 19 | 50.0308
|  | 20 | 48.3581
| Mean |  | 53.2517 | 54 | 57.3333 | 54.7948 | 51 | 52.5 | 51.5 | 54.3333 | 50.6667 | 51 | 54 | 51.8333 | 51.8333 | 46.3333 | 46.6667 | 56.611 | 54.3333 | 55.5 | 42.6267 | 51.6667 | 49.1667 | 51.3333 | 50 | 53.8333
| Median |  | 50.0703 | 52 | 57 | 54.667 | 51 | 52.1667 | 51.5 | 54 | 50.5 | 51 | 54.5 | 51.5 | 52.5 | 45 | 47.5 | 56.2616 | 53.5 | 55.5 | 41.4 | 50 | 49 | 51 | 50 | 53.5
| Std Dev. |  | 10.0778 | 5.1381 | 4.1312 | 1.3221 | 1.4142 | 3.3582 | 0.54772 | 3.6697 | 1.9664 | 1.0954 | 1.2649 | 2.0412 | 1.472 | 4.8442 | 1.9664 | 1.2619 | 3.3862 | 4.0373 | 5.6911 | 4.0825 | 2.137 | 2.4221 | 0.89443 | 1.169
| Rel.Std.Dev. |  | 18.92% | 9.515% | 7.206% | 2.413% | 2.773% | 6.397% | 1.064% | 6.754% | 3.881% | 2.148% | 2.342% | 3.938% | 2.84% | 10.46% | 4.214% | 2.229% | 6.232% | 7.274% | 13.35% | 7.902% | 4.346% | 4.718% | 1.789% | 2.172%
| PDM3 |  | 1.696% | 3.125% | 9.49% | 4.643% | -2.604% | 0.2602% | -1.65% | 3.761% | -3.241% | -2.604% | 3.125% | -1.013% | -1.013% | -11.52% | -10.88% | 8.111% | 3.761% | 5.989% | -18.6% | -1.331% | -6.106% | -1.968% | -4.514% | 2.806%
| Z-Score (Absolute) |  | NA | 0.67 | 1.71 | 0.92 | 0.26 | 0.21 | 0.1 | 0.78 | 0.36 | 0.26 | 0.67 | 0 | 0 | 1.71 | 1.61 | 1.49 | 0.78 | 1.14 | 2.87 | NA | 0.83 | 0.16 | 0.57 | 0.62
Results from laboratory 2.19 were removed due to their 10 ppb reading resolution.
| Analytical results for Pd in OREAS 906b (Indicative Value 2.17 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [22 empty] | 1
| No | No | 2.04 | 2.12 | [22 empty] | ppb
|  |  | FA*OES | FA*OES | [22 empty] | 2
|  |  | 30g | 30g | [22 empty] | 2
| 1 | 1 | < 2 | 2 | [22 empty] | 1
| 1 | 2 | < 2 | 2 | [22 empty] | 1
| 1 | 3 | < 2 | 2 | [22 empty] | 16
| 1 | 4 | < 2 | 2 | [22 empty] | 2.1667
| 1 | 5 | < 2 | 3 | [22 empty] | 7
| 1 | 6 | < 2 | 2
| Mean |  |  | 2.1667
| Median |  |  | 2
| Std Dev. |  |  | 0.40825
| Rel.Std.Dev. |  |  | 18.84%
| PDM3 |  |  | -1.665e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pt in OREAS 906b (Indicative Value < 3 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [22 empty] | 1
| No | No | 2.04 | 2.12 | [22 empty] | ppb
|  |  | FA*OES | FA*OES | [22 empty] | 2
|  |  | 30g | 30g | [22 empty] | 2
| 1 | 1 | < 3 | < 10 | [22 empty] | 1
| 1 | 2 | < 3 | < 10 | [22 empty] | 1
| 1 | 3 | < 3 | < 10 | [22 empty] | 16
| 1 | 4 | < 3 | < 10 | [22 empty] | < 3
| 1 | 5 | < 3 | < 10 | [22 empty] | 7
| 1 | 6 | < 3 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C50: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C68: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 906b (Certified Value 51.8 ppb) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.18 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | ppb
|  |  | INAA | AR*MS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*OES | AR*AAS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
|  |  | 0.085g | 25g | 30g | 25g | 15g | 50g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 25g | 50g | 15g | 30g | 15g | 30g | 2
| 1 | 1 | 67.506 | 56 | 60 | 52.0899 | 54.1 | 50 | 50 | 58 | 53 | 56 | 49 | 59 | 53.6 | 47 |  | 53.5 | 31 | 50.8 | 46.6 | 50 | 54.4 | 1
| 1 | 2 | 56.4188 | 53 | 50 | 53.8746 | 57.7 | 35 | 50 | 58 | 56 | 54 | 49 | 53 | 45 | 47 |  | 58 | 32 | 50.1 | 53.4 | 50 | 53.7 | #N/A
| 1 | 3 | 51.3899 | 49 | 50 | 53.9299 | 72.3 | 50 | 50 | 56 | 51 | 53 | 53 | 54 | 47 | 42 |  | 53.2 | 34 | 50.3 | 44.8 | 60 | 51.9 | 16
| 1 | 4 | 76.069 | 48 | 50 | 51.4769 | 69.2 | 35 | 40 | 54 | 51 | 54 | 50 | 54 | 48 | 47 |  | 55.4 | 32 | 50.2 | 44.2 | 50 | 52.4 | 51.8384
| 1 | 5 | 50.1097 | 50 | 50 | 51.6681 | 90.7 | 50 | 50 | 58 | 52 | 65 | 52 | 62 | 48.7 | 47 |  | 55.9 | 32 | 51.2 | 44.1 | 60 | 55.8 | 9
| 1 | 6 | 46.5944 | 50 | 50 | 53.7824 | 61.9 | 50 | 50 | 60 | 51 | 53 | 52 | 52 | 50.3 | 45 |  | 54.9 | 33 | 50.9 | 53.4 | 50 | 53.9
|  | 7 | 49.2208
|  | 8 | 43.0769
|  | 9 | 51.4656
|  | 10 | 47.6786
|  | 11 | 69.8264
|  | 12 | 44.6119
|  | 13 | 53.9076
|  | 14 | 48.3534
|  | 15 | 64.8784
|  | 16 | 63.8765
|  | 17 | 43.9479
|  | 18 | 37.7142
|  | 19 | 50.0308
|  | 20 | 48.3581
| Mean |  | 53.2517 | 51 | 51.6667 | 52.8036 | 67.65 | 45 | 48.3333 | 57.3333 | 52.3333 | 55.8333 | 50.8333 | 55.6667 | 48.7667 | 45.8333 |  | 55.15 | 32.3333 | 50.5833 | 47.75 | 53.3333 | 53.6833
| Median |  | 50.0703 | 50 | 50 | 52.9361 | 65.55 | 50 | 50 | 58 | 51.5 | 54 | 51 | 54 | 48.35 | 47 |  | 55.15 | 32 | 50.55 | 45.7 | 50 | 53.8
| Std Dev. |  | 10.0778 | 2.9665 | 4.0825 | 1.1775 | 13.1984 | 7.746 | 4.0825 | 2.0656 | 1.9664 | 4.6224 | 1.7224 | 3.9328 | 2.9534 | 2.0412 |  | 1.7513 | 1.0328 | 0.4446 | 4.4675 | 5.164 | 1.4049
| Rel.Std.Dev. |  | 18.92% | 5.817% | 7.902% | 2.23% | 19.51% | 17.21% | 8.447% | 3.603% | 3.757% | 8.279% | 3.388% | 7.065% | 6.056% | 4.454% |  | 3.175% | 3.194% | 0.8789% | 9.356% | 9.682% | 2.617%
| PDM3 |  | 2.727% | -1.617% | -0.3312% | 1.862% | 30.5% | -13.19% | -6.761% | 10.6% | 0.9548% | 7.707% | -1.939% | 7.385% | -5.926% | -11.58% |  | 6.388% | -37.63% | -2.421% | -7.887% | 2.884% | 3.559%
| Z-Score (Absolute) |  | NA | 0.12 | NA | 0.21 | 2.91 | 1.21 | NA | 1.03 | 0.12 | 0.76 | 0.15 | 0.73 | 0.53 | 1.06 | NA | 0.63 | 3.52 | 0.2 | 0.71 | NA | 0.37
Results from laboratories 2.02, 2.06 and 2.26 were removed due to their 10 ppb reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## SQL
| Analytical results for Cu-Res(4A) in OREAS 906b (Certified Value 0.039 wt.%) | [20 empty] | Certified
Cu-Res(4A) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.22 | 2.25 | 2.26 |  |  |  | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*AAS | SQL-Tot*Calc | 4A*AAS | 4A*AAS | 4A*AAS | 4A*OES | SQL-Tot*AAS | SQL-Tot*AAS | 4A*AAS | SQL-Tot*AAS | SQL-Tot*AAS | 4A*AAS | SQL-Tot*AAS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.0342 | 0.02 | 0.029477 | 0.039 | 0.04 | 0.04 | 0.04 | 0.031 | 0.039 | 0.035569 | 0.042 | 0.041 | 0.03619 | 0.0404 | 0.036 | 0.042 |  |  |  | 1
| 1 | 2 | 0.0346 | 0.02 | 0.029659 | 0.039 | 0.04 | 0.04 | 0.04 | 0.0289 | 0.039 | 0.03603 | 0.038 | 0.04 | 0.03739 | 0.0416 | 0.036 | 0.044 |  |  |  | #N/A
| 1 | 3 | 0.0347 | 0.02 | 0.028937 | 0.039 | 0.04 | 0.04 | 0.04 | 0.0279 | 0.039 | 0.035866 | 0.038 | 0.039 | 0.03633 | 0.0408 | 0.036 | 0.043 |  |  |  | 16
| 1 | 4 | 0.0347 | 0.02 | 0.029382 | 0.049 | 0.04 | 0.04 | 0.03 | 0.0232 | 0.039 | 0.036035 | 0.041 | 0.038 | 0.03699 | 0.0392 | 0.037 | 0.043 |  |  |  | 0.038743
| 1 | 5 | 0.0345 | 0.02 | 0.029548 | 0.039 | 0.04 | 0.04 | 0.04 | 0.0259 | 0.038 | 0.035928 | 0.04 | 0.041 | 0.03663 | 0.04 | 0.037 | 0.042 |  |  |  | 11
| 1 | 6 | 0.034 | 0.029 | 0.028984 | 0.039 | 0.04 | 0.04 | 0.04 | 0.0294 | 0.038 | 0.036138 | 0.038 | 0.041 | 0.03674 | 0.0394 | 0.036 | 0.043
| Mean |  | 0.03445 | 0.0215 | 0.029331 | 0.040667 | 0.04 | 0.04 | 0.038333 | 0.027717 | 0.038667 | 0.035928 | 0.0395 | 0.04 | 0.036712 | 0.040233 | 0.036333 | 0.042833
| Median |  | 0.03455 | 0.02 | 0.029429 | 0.039 | 0.04 | 0.04 | 0.04 | 0.0284 | 0.039 | 0.035979 | 0.039 | 0.0405 | 0.036685 | 0.0402 | 0.036 | 0.043
| Std Dev. |  | 0.0002881 | 0.0036742 | 0.00030145 | 0.0040825 | 0 | 0 | 0.0040825 | 0.0027838 | 0.0005164 | 0.00019947 | 0.0017607 | 0.0012649 | 0.00043884 | 0.00089815 | 0.0005164 | 0.00075277
| Rel.Std.Dev. |  | 0.8363% | 17.09% | 1.028% | 10.04% | 0% | 0% | 10.65% | 10.04% | 1.336% | 0.5552% | 4.457% | 3.162% | 1.195% | 2.232% | 1.421% | 1.757%
| PDM3 |  | -11.08% | -44.51% | -24.29% | 4.966% | 3.245% | 3.245% | -1.057% | -28.46% | -0.1964% | -7.266% | 1.955% | 3.245% | -5.243% | 3.847% | -6.219% | 10.56%
| Z-Score (Absolute) |  | 1.38 | 5.8 | 3.13 | 0.74 | 0.51 | 0.51 | 0.06 | 3.68 | 0.06 | 0.88 | 0.34 | 0.51 | 0.61 | 0.59 | 0.74 | 1.48
| Analytical results for Cu-Sol(CN) in OREAS 906b (Certified Value 0.015 wt.%) | [20 empty] | Certified
Cu-Sol(CN) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.22 | 2.25 | 2.26 |  |  |  | wt.%
|  |  | AL-NaCN*AAS | NaCN*AAS | NaCN*AAS | SQL-Tot*Calc | CNL*AAS | NaCN*AAS | NaCN*AAS | AL-NaCN*AAS | SQL-Tot*AAS | SQL-Tot*AAS | NaCN*AAS | SQL-Tot*AAS | SQL-Tot*AAS | AL-NaCN*AAS | SQL-Tot*AAS | CNL*OES |  |  |  | 3
3
| 1 | 1 | 0.015 | 0.0125 | 0.014802 | 0.015 | 0.02 | 0.01 | 0.01 | 0.014024 | 0.017 | 0.030785 | 0.014 | 0.016 | 0.01904 | 0.018 | 0.015 | 0.017 |  |  |  | 1
| 1 | 2 | 0.0151 | 0.0122 | 0.015505 | 0.015 | 0.02 | 0.01 | 0.01 | 0.013657 | 0.017 | 0.029824 | 0.014 | 0.018 | 0.01976 | 0.0178 | 0.015 | 0.017 |  |  |  | #N/A
| 1 | 3 | 0.0143 | 0.0125 | 0.014803 | 0.015 | 0.02 | 0.01 | 0.01 | 0.014688 | 0.017 | 0.029373 | 0.014 | 0.016 | 0.01961 | 0.0182 | 0.015 | 0.015 |  |  |  | 16
| 1 | 4 | 0.0145 | 0.0121 | 0.0152 | 0.016 | 0.02 | 0.02 | 0.01 | 0.013674 | 0.018 | 0.029236 | 0.014 | 0.019 | 0.01974 | 0.017 | 0.015 | 0.018 |  |  |  | 0.0154
| 1 | 5 | 0.015 | 0.0125 | 0.014866 | 0.014 | 0.02 | 0.01 | 0.01 | 0.013726 | 0.018 | 0.031783 | 0.014 | 0.018 | 0.01949 | 0.0166 | 0.016 | 0.019 |  |  |  | 12
| 1 | 6 | 0.0151 | 0.0123 | 0.015536 | 0.015 | 0.02 | 0.01 | 0.01 | 0.013591 | 0.017 | 0.033119 | 0.014 | 0.016 | 0.01961 | 0.017 | 0.015 | 0.019
| Mean |  | 0.014833 | 0.01235 | 0.015118 | 0.015 | 0.02 | 0.011667 | 0.01 | 0.013893 | 0.017333 | 0.030687 | 0.014 | 0.017167 | 0.019542 | 0.017433 | 0.015167 | 0.0175
| Median |  | 0.015 | 0.0124 | 0.015033 | 0.015 | 0.02 | 0.01 | 0.01 | 0.0137 | 0.017 | 0.030304 | 0.014 | 0.017 | 0.01961 | 0.0174 | 0.015 | 0.0175
| Std Dev. |  | 0.00034448 | 0.00017607 | 0.00034445 | 0.00063246 | 0 | 0.0040825 | 0 | 0.0004176 | 0.0005164 | 0.0015299 | 1.9003e-18 | 0.0013292 | 0.00026483 | 0.00065013 | 0.00040825 | 0.0015166
| Rel.Std.Dev. |  | 2.322% | 1.426% | 2.278% | 4.216% | 0% | 34.99% | 0% | 3.006% | 2.979% | 4.986% | 1.357e-14% | 7.743% | 1.355% | 3.729% | 2.692% | 8.666%
| PDM3 |  | -3.681% | -19.81% | -1.83% | -2.599% | 29.87% | -24.24% | -35.07% | -9.785% | 12.55% | 99.26% | -9.092% | 11.47% | 26.89% | 13.2% | -1.517% | 13.63%
| Z-Score (Absolute) |  | 0.09 | 0.84 | 0.01 | 0.04 | 1.46 | 1.05 | 1.55 | 0.38 | 0.66 | 4.68 | 0.34 | 0.61 | 1.32 | 0.69 | 0.01 | 0.71
| Analytical results for Cu-Sol(H2SO4) in OREAS 906b (Certified Value 0.26 wt.%) | [20 empty] | Certified
Cu-Sol(H2SO4) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | SQL-Tot*Calc | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | SQL-Tot*AAS | SQL-Tot*AAS | AL-5S*AAS | SQL-Tot*AAS | AL-5S*OES | SQL-Tot*AAS | AL-5S*AAS | AL-5S*AAS | SQL-Tot*AAS | AL-5S*OES/AAS | AL-5S*AAS | 3
3
| 1 | 1 | 0.2513 | 0.273 | 0.26112 | 0.254 | 0.25 | 0.27 | 0.27 | 0.27 | 0.262 | 0.24065 | 0.266 | 0.263 | 0.26 | 0.25239 | 0.2677 | 0.2828 | 0.244 | 0.258 | 0.265 | 1
| 1 | 2 | 0.2518 | 0.272 | 0.25798 | 0.255 | 0.26 | 0.27 | 0.26 | 0.26 | 0.261 | 0.24351 | 0.263 | 0.263 | 0.25 | 0.25234 | 0.26387 | 0.2804 | 0.252 | 0.258 | 0.265 | #N/A
| 1 | 3 | 0.2585 | 0.273 | 0.25443 | 0.253 | 0.26 | 0.27 | 0.26 | 0.26 | 0.262 | 0.256 | 0.267 | 0.261 | 0.25727 | 0.25369 | 0.26617 | 0.2832 | 0.251 | 0.263 | 0.264 | 16
| 1 | 4 | 0.2571 | 0.278 | 0.25539 | 0.255 | 0.26 | 0.26 | 0.26 | 0.27 | 0.262 | 0.24864 | 0.266 | 0.261 | 0.25 | 0.25428 | 0.26117 | 0.282 | 0.247 | 0.26 | 0.262 | 0.25975
| 1 | 5 | 0.2531 | 0.264 | 0.26148 | 0.255 | 0.26 | 0.27 | 0.26 | 0.26 | 0.26 | 0.25091 | 0.267 | 0.264 | 0.25979 | 0.2545 | 0.26375 | 0.2834 | 0.249 | 0.257 | 0.27 | 13
| 1 | 6 | 0.2567 | 0.279 | 0.25714 | 0.253 | 0.26 | 0.27 | 0.26 | 0.26 | 0.261 | 0.24537 | 0.267 | 0.266 | 0.26 | 0.25294 | 0.26883 | 0.2828 | 0.241 | 0.255 | 0.266
| Mean |  | 0.25475 | 0.27317 | 0.25792 | 0.25417 | 0.25833 | 0.26833 | 0.26167 | 0.26333 | 0.26133 | 0.24751 | 0.266 | 0.263 | 0.25618 | 0.25336 | 0.26525 | 0.28243 | 0.24733 | 0.2585 | 0.26533
| Median |  | 0.2549 | 0.273 | 0.25756 | 0.2545 | 0.26 | 0.27 | 0.26 | 0.26 | 0.2615 | 0.247 | 0.2665 | 0.263 | 0.25853 | 0.25331 | 0.26502 | 0.2828 | 0.248 | 0.258 | 0.265
| Std Dev. |  | 0.0030566 | 0.0053448 | 0.0029038 | 0.00098319 | 0.0040825 | 0.0040825 | 0.0040825 | 0.005164 | 0.0008165 | 0.0055213 | 0.0015492 | 0.0018974 | 0.0048945 | 0.00093942 | 0.0028438 | 0.0011057 | 0.0042269 | 0.0027386 | 0.0026583
| Rel.Std.Dev. |  | 1.2% | 1.957% | 1.126% | 0.3868% | 1.58% | 1.521% | 1.56% | 1.961% | 0.3124% | 2.231% | 0.5824% | 0.7214% | 1.911% | 0.3708% | 1.072% | 0.3915% | 1.709% | 1.059% | 1.002%
| PDM3 |  | -1.924% | 5.166% | -0.7023% | -2.149% | -0.5447% | 3.305% | 0.7385% | 1.38% | 0.6102% | -4.711% | 2.407% | 1.252% | -1.375% | -2.461% | 2.118% | 8.733% | -4.78% | -0.4806% | 2.15%
| Z-Score (Absolute) |  | 0.95 | 1.71 | 0.49 | 1.04 | 0.43 | 1.01 | 0.05 | 0.29 | 0 | 2 | 0.67 | 0.24 | 0.75 | 1.15 | 0.57 | 3.05 | 2.02 | 0.41 | 0.58
| Analytical results for Cu-Tot(Calc) in OREAS 906b (Certified Value 0.313 wt.%) | [20 empty] | Certified
Cu-Tot(Calc) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.22 | 2.25 | 2.26 |  |  |  | wt.%
|  |  | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | SQL-Tot*Calc | 4A*AAS | SQL-Tot*Calc | SQL-Tot*AAS | SQL-Tot*Calc | SQL-Tot*AAS | SQL-Tot*Calc |  |  |  | 3
3
| 1 | 1 | 0.3005 | 0.3055 | 0.3054 | 0.309 | 0.31 | 0.32 | 0.32 | 0.31502 | 0.318 | 0.30701 | 0.321 | 0.32 | 0.30761 | 0.3412 | 0.31 | 0.317 |  |  |  | 1
| 1 | 2 | 0.3015 | 0.3042 | 0.30314 | 0.308 | 0.32 | 0.32 | 0.31 | 0.30256 | 0.317 | 0.30937 | 0.315 | 0.321 | 0.3095 | 0.3398 | 0.318 | 0.319 |  |  |  | #N/A
| 1 | 3 | 0.3075 | 0.3055 | 0.29817 | 0.307 | 0.32 | 0.32 | 0.31 | 0.30259 | 0.318 | 0.32123 | 0.319 | 0.316 | 0.30962 | 0.3422 | 0.31 | 0.321 |  |  |  | 16
| 1 | 4 | 0.3063 | 0.3101 | 0.29997 | 0.32 | 0.32 | 0.32 | 0.3 | 0.30687 | 0.319 | 0.31391 | 0.321 | 0.318 | 0.31101 | 0.3382 | 0.311 | 0.321 |  |  |  | 0.31258
| 1 | 5 | 0.3026 | 0.2965 | 0.3059 | 0.309 | 0.32 | 0.32 | 0.31 | 0.29963 | 0.316 | 0.31862 | 0.321 | 0.323 | 0.31062 | 0.34 | 0.316 | 0.318 |  |  |  | 14
| 1 | 6 | 0.3058 | 0.3203 | 0.30166 | 0.307 | 0.32 | 0.32 | 0.31 | 0.30299 | 0.316 | 0.31462 | 0.318 | 0.323 | 0.30928 | 0.3392 | 0.312 | 0.317
| Mean |  | 0.30403 | 0.30702 | 0.30237 | 0.31 | 0.31833 | 0.32 | 0.31 | 0.30494 | 0.31733 | 0.31413 | 0.31917 | 0.32017 | 0.30961 | 0.3401 | 0.31283 | 0.31883
| Median |  | 0.3042 | 0.3055 | 0.3024 | 0.3085 | 0.32 | 0.32 | 0.31 | 0.30279 | 0.3175 | 0.31427 | 0.32 | 0.3205 | 0.30956 | 0.3399 | 0.3115 | 0.3185
| Std Dev. |  | 0.0028717 | 0.0078624 | 0.0030378 | 0.00498 | 0.0040825 | 0 | 0.0063246 | 0.005453 | 0.0012111 | 0.0053728 | 0.0024014 | 0.0027869 | 0.0011917 | 0.0014241 | 0.0033714 | 0.0018348
| Rel.Std.Dev. |  | 0.9445% | 2.561% | 1.005% | 1.606% | 1.282% | 0% | 2.04% | 1.788% | 0.3816% | 1.71% | 0.7524% | 0.8704% | 0.3849% | 0.4187% | 1.078% | 0.5755%
| PDM3 |  | -2.736% | -1.781% | -3.267% | -0.8268% | 1.839% | 2.372% | -0.8268% | -2.444% | 1.519% | 0.4934% | 2.106% | 2.426% | -0.9526% | 8.803% | 0.07962% | 1.999%
| Z-Score (Absolute) |  | 1.15 | 0.79 | 1.36 | 0.43 | 0.59 | 0.8 | 0.43 | 1.04 | 0.47 | 0.08 | 0.69 | 0.82 | 0.47 | 3.25 | 0.08 | 0.65
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Pycnometry
| Analytical results for SG in OREAS 906b (Certified Value 2.74 Unity) | [21 empty] | Certified
SG | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.21 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 | Unity
|  |  | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | LIQPYC | LIQPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | 2
3
| 1 | 1 | 2.75 | 2.7 | 2.7692 | 2.77 | 2.7379 | 2.82 | 2.63 | 2.59 | 2.79 | 2.7 | 2.7465 | 2.55 | 2.6501 | 2.75 | 2.76 | 2.73 | 2.69 | 2.77 | 2.8 | 2.77 | 1
| 1 | 2 | 2.75 | 2.7 | 2.7894 | 2.77 | 2.7312 | 2.81 | 2.71 | 2.67 | 2.76 | 2.72 | 2.7738 | 2.59 | 2.6542 | 2.75 | 2.76 | 2.71 | 2.63 | 2.78 | 2.79 | 2.76 | #N/A
| 1 | 3 | 2.75 | 2.69 | 2.7894 | 2.78 | 2.7558 | 2.82 | 2.65 | 2.65 | 2.77 | 2.71 | 2.7792 | 2.54 | 2.5157 | 2.76 | 2.76 | 2.73 | 2.64 | 2.78 | 2.77 | 2.76 | 16
| 1 | 4 | 2.76 | 2.71 | 2.7692 | 2.78 | 2.7292 | 2.82 | 2.63 | 2.65 | 2.77 | 2.71 | 2.7611 | 2.6 | 2.5681 | 2.75 | 2.75 | 2.71 | 2.61 | 2.78 | 2.81 | 2.76 | 2.7413
| 1 | 5 | 2.76 | 2.7 | 2.8 | 2.79 | 2.7496 | 2.82 | 2.69 | 2.56 | 2.79 | 2.7 | 2.7612 | 2.6 | 2.6661 | 2.75 | 2.76 | 2.69 | 2.67 | 2.77 | 2.78 | 2.76 | 16
| 1 | 6 | 2.76 | 2.7 | 2.7592 | 2.77 | 2.7381 | 2.8 | 2.68 | 2.64 | 2.78 | 2.69 | 2.7593 | 2.55 | 2.6668 | 2.75 | 2.75 | 2.73 | 2.7 | 2.77 | 2.78 | 2.76
| Mean |  | 2.755 | 2.7 | 2.7794 | 2.7767 | 2.7403 | 2.815 | 2.665 | 2.6267 | 2.7767 | 2.705 | 2.7635 | 2.5717 | 2.6202 | 2.7517 | 2.7567 | 2.7167 | 2.6567 | 2.775 | 2.7883 | 2.7617
| Median |  | 2.755 | 2.7 | 2.7793 | 2.775 | 2.738 | 2.82 | 2.665 | 2.645 | 2.775 | 2.705 | 2.7612 | 2.57 | 2.6522 | 2.75 | 2.76 | 2.72 | 2.655 | 2.775 | 2.785 | 2.76
| Std Dev. |  | 0.0054772 | 0.0063246 | 0.015741 | 0.008165 | 0.010425 | 0.0083666 | 0.033317 | 0.042269 | 0.012111 | 0.010488 | 0.011579 | 0.027869 | 0.06321 | 0.0040825 | 0.005164 | 0.01633 | 0.03559 | 0.0054772 | 0.01472 | 0.0040825
| Rel.Std.Dev. |  | 0.1988% | 0.2342% | 0.5663% | 0.2941% | 0.3804% | 0.2972% | 1.25% | 1.609% | 0.4362% | 0.3877% | 0.419% | 1.084% | 2.412% | 0.1484% | 0.1873% | 0.6011% | 1.34% | 0.1974% | 0.5279% | 0.1478%
| PDM3 |  | 0.5014% | -1.505% | 1.392% | 1.292% | -0.0348% | 2.69% | -2.782% | -4.18% | 1.292% | -1.323% | 0.8125% | -6.186% | -4.416% | 0.3798% | 0.5622% | -0.8969% | -3.086% | 1.231% | 1.717% | 0.7446%
| Z-Score (Absolute) |  | 0.04 | 1.18 | 0.58 | 0.52 | 0.29 | 1.36 | 1.95 | 2.8 | 0.52 | 1.07 | 0.23 | 4.01 | 2.94 | 0.04 | 0.07 | 0.81 | 2.13 | 0.48 | 0.77 | 0.18
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 906b (Indicative Value 0.124 wt.%) | [9 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.06 | 2.22 | 3.01 | [5 empty] | wt.%
|  |  | IRC | IRC-x*Calc | IRC | [5 empty] | 3
3
| 1 | 1 | 0.12 |  | 0.11 | [5 empty] | 1
| 1 | 2 | 0.13 |  | 0.14 | [5 empty] | 3
| 1 | 3 | 0.12 | [7 empty] | 16
| 1 | 4 | 0.12 | [7 empty] | 0.12417
| 1 | 5 | 0.12 | [7 empty] | 9
| 1 | 6 | 0.13
| Mean |  | 0.12333 |  | 0.125
| Median |  | 0.12 |  | 0.125
| Std Dev. |  | 0.005164 |  | 0.021213
| Rel.Std.Dev. |  | 4.187% |  | 16.97%
| PDM3 |  | -0.6711% |  | 0.6711%
| Z-Score (Absolute) |  | 0.67 | NA | 0.67
| Analytical results for S in OREAS 906b (Indicative Value 0.06 wt.%) | [9 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 | 3.01 | wt.%
|  |  | IRC | IRC-x*Calc | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | NA | - | 0.05 | -- |  | N.A. | N.A. | 0.06 | 1
| 1 | 2 | NA | - | 0.05 | -- |  | N.A. | N.A. | 0.08 | 3
| 1 | 3 | NA | - | 0.05 | -- |  | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 0.05 | -- |  | N.A. | N.A. |  | 0.06
| 1 | 5 | NA | - | 0.05 | -- |  | N.A. | N.A. |  | 9
| 1 | 6 | NA | - | 0.05 | -- |  | N.A. | N.A.
| Mean |  |  |  | 0.05 |  |  |  |  | 0.07
| Median |  |  |  | 0.05 |  |  |  |  | 0.07
| Std Dev. |  |  |  | 7.6012e-18 |  |  |  |  | 0.014142
| Rel.Std.Dev. |  |  |  | 1.52e-14% |  |  |  |  | 20.2%
| PDM3 |  |  |  | -16.67% |  |  |  |  | 16.67%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | NA | NA | NA | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ALK
| Analytical results for S-(Sulphate) in OREAS 906b (Indicative Value 0.045 wt.%) | [8 empty] | Indicative
S-(Sulphate) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | Na2CO3*GRAV | 3
3
| 1 | 1 | NA | - | 0.5 | -- |  | N.A. | N.A. | 1
| 1 | 2 | NA | - | 0.05 | -- |  | N.A. | N.A. | 5
| 1 | 3 | NA | - | 0.05 | -- |  | N.A. | N.A. | 16
| 1 | 4 | NA | - | < 0.05 | -- |  | N.A. | N.A. | 0.045
| 1 | 5 | NA | - | 0.05 | -- |  | N.A. | N.A. | 11
| 1 | 6 | NA | - | 0.05 | -- |  | N.A. | N.A.
| Mean |  |  |  | 0.14
| Median |  |  |  | 0.05
| Std Dev. |  |  |  | 0.20125
| Rel.Std.Dev. |  |  |  | 143.7%
| PDM3 |  |  |  | 211.1%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for S-(Sulphide) in OREAS 906b (Indicative Value < 0.05 wt.%) | [8 empty] | Indicative
S-(Sulphide) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 |  | wt.%
|  |  | IRC-x*Calc | Na2CO3*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC-x*Calc |  | 3
3
| 1 | 1 | - | < 0.05 | -- |  | N.A. | N.A. |  | 1
| 1 | 2 | - | < 0.05 | -- |  | N.A. | N.A. |  | 5
| 1 | 3 | - | < 0.05 | -- |  | N.A. | N.A. |  | 16
| 1 | 4 | - | < 0.05 | -- |  | N.A. | N.A. |  | < 0.05
| 1 | 5 | - | < 0.05 | -- |  | N.A. | N.A. |  | 11
| 1 | 6 | - | < 0.05 | -- |  | N.A. | N.A.
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 906b (Certified Value 0.709 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.8 | 0.7 | 0.75082 | 0.681 | 0.6 | 0.68 | 0.58 | 0.7 | 0.68 | 0.72 | 0.72 | 1.04 | 0.72 | 0.72535 | < 2 | 0.9 |  | < 0.5 | 0.7 | 0.7 | 0.688 | 0.76 | 1.4 | 0.7 | 0.712 | 1
| 1 | 2 | 0.76 | 0.7 | 0.75153 | 0.683 | 0.61 | 0.66 | 0.54 | 0.68 | 0.74 | 0.65 | 0.72 | 0.98 | 0.66 | 0.77211 | < 2 | 0.9 |  | < 0.5 | 0.67 | 0.7 | 0.667 | 0.76 | 1.1 | 0.7 | 0.663 | 9
| 1 | 3 | 0.73 | 0.7 | 0.73555 | 0.68 | 0.61 | 0.73 | 0.53 | 1.68 | 0.7 | 0.68 | 0.72 | 1.09 | 0.69 | 0.72536 | < 2 | 0.8 |  | < 0.5 | 0.69 | 0.8 | 0.669 | 0.75 | 1.3 | 0.6 | 0.684 | 16
| 1 | 4 | 0.75 | 0.7 | 0.73312 | 0.697 | 0.6 | 0.71 | 0.49 | 0.7 | 0.75 | 0.68 | 0.68 | 1 | 0.69 | 0.72318 | < 2 | 0.9 |  | < 0.5 | 0.72 | 0.8 | 0.682 | 0.76 | 1.3 | 0.7 | 0.631 | 0.70937
| 1 | 5 | 0.72 | 0.7 | 0.74901 | 0.668 | 0.61 | 0.74 | 0.53 | 0.68 | 0.75 | 0.65 | 0.69 | 1.24 | 0.69 | 0.75843 | < 2 | 0.9 |  | < 0.5 | 0.72 | 0.8 | 0.684 | 0.76 | 1.4 | 0.7 | 0.683 | 18
| 1 | 6 | 0.75 | 0.7 | 0.74169 | 0.716 | 0.61 | 0.72 | 0.5 | 0.7 | 0.69 | 0.7 | 0.73 | 1.03 | 0.66 | 0.79512 | < 2 | 0.9 |  | < 0.5 | 0.68 | 0.8 | 0.679 | 0.76 | 1.4 | 0.7 | 0.667
| Mean |  | 0.75167 | 0.7 | 0.74362 | 0.6875 | 0.60667 | 0.70667 | 0.52833 | 0.85667 | 0.71833 | 0.68 | 0.71 | 1.0633 | 0.685 | 0.74992 |  | 0.88333 |  |  | 0.69667 | 0.76667 | 0.67817 | 0.75833 | 1.3167 | 0.68333 | 0.67333
| Median |  | 0.75 | 0.7 | 0.74535 | 0.682 | 0.61 | 0.715 | 0.53 | 0.7 | 0.72 | 0.68 | 0.72 | 1.035 | 0.69 | 0.74189 |  | 0.9 |  |  | 0.695 | 0.8 | 0.6805 | 0.76 | 1.35 | 0.7 | 0.675
| Std Dev. |  | 0.027869 | 1.2162e-16 | 0.0080326 | 0.016742 | 0.005164 | 0.030768 | 0.031885 | 0.40347 | 0.031885 | 0.027568 | 0.02 | 0.094375 | 0.022583 | 0.030102 |  | 0.040825 |  |  | 0.020656 | 0.05164 | 0.0084242 | 0.0040825 | 0.1169 | 0.040825 | 0.026986
| Rel.Std.Dev. |  | 3.708% | 1.737e-14% | 1.08% | 2.435% | 0.8512% | 4.354% | 6.035% | 47.1% | 4.439% | 4.054% | 2.817% | 8.875% | 3.297% | 4.014% |  | 4.622% |  |  | 2.965% | 6.736% | 1.242% | 0.5383% | 8.879% | 5.974% | 4.008%
| PDM3 |  | 5.962% | -1.321% | 4.828% | -3.083% | -14.48% | -0.3814% | -25.52% | 20.76% | 1.263% | -4.141% | 0.08852% | 49.9% | -3.436% | 5.717% |  | 24.52% |  |  | -1.791% | 8.077% | -4.399% | 6.902% | 85.61% | -3.671% | -5.08%
| Z-Score (Absolute) |  | 0.82 | NA | 0.67 | 0.35 | 1.82 | 0 | 3.25 | 2.74 | 0.21 | 0.49 | 0.06 | 6.51 | 0.4 | 0.79 | 5.35 | NA | NA | 8.33 | 0.18 | NA | 0.52 | 0.94 | NA | NA | 0.61
Results from laboratories 2.02, 2.17, 2.22, 2.25 and 2.26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 906b (Certified Value 7.46 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 7.5563 | 7.09 | 7.418 | 7.5 | 8 | 7.48 | 7.23 | 7.5 | 7.75 | 7.64 | 7.26 | 8.12 | 7.66 | 7.5547 | 7.47 | 7.2 | 7.6597 | 7.23 | 7.5 | 7.36 | 7.5 | 6.35 | 6.53 | 7.1062 | 7.87 | 1
| 1 | 2 | 7.526 | 7.12 | 7.2717 | 7.73 | 7.9 | 7.54 | 6.94 | 7.66 | 7.75 | 7.52 | 7.23 | 7.51 | 7.7 | 7.5407 | 7.6 | 7.21 | 7.6598 | 7.46 | 6.74 | 7.49 | 7.62 | 6.31 | 7.62 | 7.0845 | 7.75 | #N/A
| 1 | 3 | 7.4989 | 7.12 | 7.2772 | 7.58 | 7.83 | 7.54 | 7.11 | 7.68 | 7.72 | 7.21 | 7.28 | 7.82 | 7.86 | 7.6565 | 7.34 | 7.21 | 7.6562 | 7.37 | 7.33 | 7.07 | 7.52 | 6.35 | 7.44 | 7.0632 | 7.45 | 16
| 1 | 4 | 7.4519 | 7.16 | 7.4184 | 7.56 | 7.9 | 7.49 | 7.07 | 7.71 | 7.82 | 7.39 | 7.16 | 7.72 | 7.46 | 7.4834 | 7.4 | 7.16 | 7.6508 | 7.45 | 6.85 | 7.51 | 7.4473 | 6.36 | 7.23 | 7.0705 | 7.62 | 7.4554
| 1 | 5 | 7.5295 | 7.26 | 7.3639 | 7.61 | 7.89 | 7.64 | 7.13 | 7.65 | 7.73 | 7.66 | 7.26 | 8.13 | 7.55 | 7.4595 | 7.31 | 7.19 | 7.6579 | 7.44 | 7.18 | 7.43 | 7.58 | 6.38 | 6.51 | 7.0276 | 7.49 | 19
| 1 | 6 | 7.5093 | 7.01 | 7.294 | 7.54 | 7.79 | 7.33 | 7 | 7.64 | 7.59 | 7.51 | 7.23 | 7.94 | 7.72 | 7.6368 | 7.36 | 7.18 | 7.6529 | 7.41 | 6.92 | 6.46 | 7.5 | 6.34 | 7.34 | 7.007 | 7.47
| Mean |  | 7.512 | 7.1267 | 7.3405 | 7.5867 | 7.885 | 7.5033 | 7.08 | 7.64 | 7.7267 | 7.4883 | 7.2367 | 7.8733 | 7.6583 | 7.5553 | 7.4133 | 7.1917 | 7.6562 | 7.3933 | 7.0867 | 7.22 | 7.5279 | 6.3483 | 7.1117 | 7.0598 | 7.6083
| Median |  | 7.5176 | 7.12 | 7.3289 | 7.57 | 7.895 | 7.515 | 7.09 | 7.655 | 7.74 | 7.515 | 7.245 | 7.88 | 7.68 | 7.5477 | 7.38 | 7.195 | 7.657 | 7.425 | 7.05 | 7.395 | 7.51 | 6.35 | 7.285 | 7.0669 | 7.555
| Std Dev. |  | 0.035371 | 0.082381 | 0.068601 | 0.079415 | 0.071764 | 0.10211 | 0.10198 | 0.072938 | 0.075542 | 0.16798 | 0.042269 | 0.24064 | 0.13949 | 0.079319 | 0.1069 | 0.019408 | 0.0036918 | 0.08641 | 0.29757 | 0.40517 | 0.062149 | 0.023166 | 0.47596 | 0.036644 | 0.1714
| Rel.Std.Dev. |  | 0.4709% | 1.156% | 0.9346% | 1.047% | 0.9101% | 1.361% | 1.44% | 0.9547% | 0.9777% | 2.243% | 0.5841% | 3.056% | 1.821% | 1.05% | 1.442% | 0.2699% | 0.04822% | 1.169% | 4.199% | 5.612% | 0.8256% | 0.3649% | 6.693% | 0.519% | 2.253%
| PDM3 |  | 0.759% | -4.409% | -1.541% | 1.761% | 5.762% | 0.643% | -5.035% | 2.476% | 3.639% | 0.4418% | -2.934% | 5.606% | 2.722% | 1.34% | -0.5642% | -3.537% | 2.694% | -0.8325% | -4.946% | -3.157% | 0.9723% | -14.85% | -4.61% | -5.306% | 2.051%
| Z-Score (Absolute) |  | 0.09 | 1.43 | 0.59 | 0.39 | 1.57 | 0.06 | 1.62 | 0.6 | 0.95 | 0 | 1 | 1.53 | 0.67 | 0.27 | 0.3 | 1.18 | 0.67 | 0.38 | 1.59 | 1.06 | 0.16 | 4.52 | 1.49 | 1.7 | 0.48
| Analytical results for As in OREAS 906b (Certified Value 13.7 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 13.3 | 13 | 13.3324 | 14.3 | 2.98 | 14.1 | 13.8 | 13.6 | 14 | 11.5 | 13 | 18.3 | 15 | 16.2402 | 14 | 14.9 | 13 | 12.3 | 12.2 | 21 | 9 | 16.5 | 12.9 | 13 |  | 1
| 1 | 2 | 13.2 | 13 | 13.3559 | 14.4 | 2.74 | 13.9 | 13.3 | 13.2 | 13.6 | 12 | 13 | 18.1 | 14 | 15.9534 | 15 | 14.8 | 13 | 12 | 13 | 22 | 9 | 14.7 | 12.9 | 12.6 |  | 10
| 1 | 3 | 12.8 | 13 | 13.1329 | 15.2 | 3.33 | 14.5 | 13.9 | 13.7 | 14.3 | 13 | 12 | 22.7 | 14 | 16.3071 | 11 | 14.5 | 13 | 12.6 | 13.7 | 21 | 9 | 16.1 | 12.7 | 12.4 |  | 16
| 1 | 4 | 13.2 | 13 | 13.3269 | 15 | 2.94 | 13.6 | 13.3 | 13.1 | 14.4 | 11.6 | 12 | 16.8 | 14 | 16.001 | 12 | 14.5 | 13 | 12.4 | 12 | 23 | 9 | 17.1 | 13.4 | 11.8 |  | 13.6874
| 1 | 5 | 13.3 | 13 | 13.1173 | 14.1 | 2.48 | 13.8 | 13.8 | 13.1 | 15.1 | 12.2 | 11 | 15.9 | 15 | 15.9956 | 13 | 14.3 | 13 | 12.6 | 12.6 | 23 | 9 | 16.6 | 13.7 | 12.8 |  | 20
| 1 | 6 | 13.1 | 13 | 13.2623 | 13.8 | 3.41 | 13.9 | 13.5 | 13 | 13.7 | 12.6 | 13 | 16.6 | 14 | 15.9215 | 11 | 13.8 | 12 | 12.1 | 12.6 | 20 | 9 | 16.1 | 12.7 | 11.8
| Mean |  | 13.15 | 13 | 13.2546 | 14.4667 | 2.98 | 13.9667 | 13.6 | 13.2833 | 14.1833 | 12.15 | 12.3333 | 18.0667 | 14.3333 | 16.0698 | 12.6667 | 14.4667 | 12.8333 | 12.3333 | 12.6833 | 21.6667 | 9 | 16.1833 | 13.05 | 12.4
| Median |  | 13.2 | 13 | 13.2946 | 14.35 | 2.96 | 13.9 | 13.65 | 13.15 | 14.15 | 12.1 | 12.5 | 17.45 | 14 | 15.9983 | 12.5 | 14.5 | 13 | 12.35 | 12.6 | 21.5 | 9 | 16.3 | 12.9 | 12.5
| Std Dev. |  | 0.18708 | 0 | 0.10516 | 0.53541 | 0.35117 | 0.30768 | 0.26833 | 0.29269 | 0.54924 | 0.57879 | 0.8165 | 2.4484 | 0.5164 | 0.16194 | 1.633 | 0.39328 | 0.40825 | 0.25033 | 0.608 | 1.2111 | 0 | 0.81588 | 0.40866 | 0.50596
| Rel.Std.Dev. |  | 1.423% | 0% | 0.7934% | 3.701% | 11.78% | 2.203% | 1.973% | 2.203% | 3.872% | 4.764% | 6.62% | 13.55% | 3.603% | 1.008% | 12.89% | 2.719% | 3.181% | 2.03% | 4.794% | 5.59% | 0% | 5.042% | 3.131% | 4.08%
| PDM3 |  | -3.926% | -5.022% | -3.162% | 5.693% | -78.23% | 2.04% | -0.6386% | -2.952% | 3.623% | -11.23% | -9.893% | 31.99% | 4.719% | 17.41% | -7.458% | 5.693% | -6.24% | -9.893% | -7.336% | 58.3% | -34.25% | 18.24% | -4.657% | -9.406%
| Z-Score (Absolute) |  | 0.1 | NA | 0.02 | 0.89 | 7.72 | 0.51 | 0.24 | 0 | 0.67 | 0.85 | NA | 3.58 | NA | 2.09 | NA | 0.89 | NA | 0.71 | 0.45 | NA | NA | 2.17 | 0.17 | 0.66
Results from laboratories 2.02, 2.12, 2.14, 2.16, 2.19, 2.23 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 906b (Indicative Value < 10 ppm) | [26 empty] | Indicative
B | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.22 | [24 empty] | ppm
|  |  | 4A*MS | [24 empty] | 1
1
| 1 | 1 | < 10 | [24 empty] | 1
| 1 | 2 | < 10 | [24 empty] | 7
| 1 | 3 | < 10 | [24 empty] | 16
| 1 | 4 | < 10 | [24 empty] | < 10
| 1 | 5 | < 10 | [24 empty] | 13
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 906b (Certified Value 2819 ppm) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 0
0
| 1 | 1 | 2894.9 | 2800 | 2824.3383 | 2702 | 3173.1 | 2800 | 2770 | 2880 | 2940 | 2620 | 2776 | 2570 | 2723 | 2794.7531 | 2843 | > 2000 | 3300 | 2788 | 2716 | 2567 | 2860 | 2961 |  |  |  | 1
| 1 | 2 | 2847.2 | 2810 | 2922.5534 | 2758 | 3199.12 | 2820 | 2730 | 3000 | 2960 | 2590 | 2759 | 2510 | 2732 | 2796.0368 | 2869 | > 2000 | 3230 | 2845 | 2776 | 2574 | 2710 | 2947 |  |  |  | 35
| 1 | 3 | 2811 | 2830 | 2799.9416 | 2748 | 3243.16 | 2830 | 2720 | 3020 | 2960 | 2770 | 2770 | 2440 | 2794 | 2781.4207 | 2800 | > 2000 | 3160 | 2924 | 2810 | 2572 | 2740 | 2769 |  |  |  | 16
| 1 | 4 | 2858.1 | 2860 | 2827.0503 | 2723 | 3017.93 | 2790 | 2710 | 3010 | 2980 | 2560 | 2730 | 2430 | 2775 | 2763.1118 | 2819 | > 2000 | 3220 | 2645 | 2831 | 2573 | 2950 | 2915 |  |  |  | 2819.2182
| 1 | 5 | 2849.9 | 2830 | 2879.8619 | 2747 | 3209.8 | 2870 | 2760 | 2930 | 2950 | 2640 | 2770 | 2500 | 2708 | 2740.6395 | 2782 | > 2000 | 3230 | 2824 | 2728 | 2520 | 2790 | 2910 |  |  |  | 21
| 1 | 6 | 2879.8 | 2800 | 2920.4051 | 2710 | 3172.66 | 2730 | 2730 | 3000 | 2880 | 2580 | 2773 | 2510 | 2786 | 2795.9313 | 2792 | > 2000 | 3210 | 2591 | 2808 | 2597 | 2800 | 2819
| Mean |  | 2856.8167 | 2821.6667 | 2862.3584 | 2731.3333 | 3169.295 | 2806.6667 | 2736.6667 | 2973.3333 | 2945 | 2626.6667 | 2763 | 2493.3333 | 2753 | 2778.6489 | 2817.5 |  | 3225 | 2769.5 | 2778.1667 | 2567.1667 | 2808.3333 | 2886.8333
| Median |  | 2854 | 2820 | 2853.4561 | 2735 | 3186.11 | 2810 | 2730 | 3000 | 2955 | 2605 | 2770 | 2505 | 2753.5 | 2788.0869 | 2809.5 |  | 3225 | 2806 | 2792 | 2572.5 | 2795 | 2912.5
| Std Dev. |  | 29.0507 | 23.1661 | 52.7016 | 22.8881 | 78.6427 | 46.7618 | 23.3809 | 55.7375 | 34.4964 | 75.8068 | 17.1581 | 51.6398 | 36.3868 | 22.6016 | 33.231 |  | 45.0555 | 126.685 | 47.0762 | 25.3726 | 86.5833 | 76.0747
| Rel.Std.Dev. |  | 1.017% | 0.821% | 1.841% | 0.838% | 2.481% | 1.666% | 0.8544% | 1.875% | 1.171% | 2.886% | 0.621% | 2.071% | 1.322% | 0.8134% | 1.179% |  | 1.397% | 4.574% | 1.695% | 0.9883% | 3.083% | 2.635%
| PDM3 |  | 1.334% | 0.08685% | 1.53% | -3.117% | 12.42% | -0.4452% | -2.928% | 5.467% | 4.462% | -6.83% | -1.994% | -11.56% | -2.349% | -1.439% | -0.06094% |  | 14.39% | -1.764% | -1.456% | -8.94% | -0.3861% | 2.398%
| Z-Score (Absolute) |  | 0.61 | 0.18 | 0.67 | 0.91 | 4.39 | 0 | 0.85 | 2.02 | 1.67 | 2.18 | 0.53 | 3.79 | 0.65 | 0.34 | 0.13 | NA | 5.07 | 0.45 | 0.35 | 2.9 | 0.02 | 0.97
| Analytical results for Be in OREAS 906b (Certified Value 2.99 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 2.95 | 2.9 | 2.9335 | 4 | 1.94 | 2.9 | 3 | 2.97 | 3.01 | 2.9 | 3.04 | 3.25 | 3 | 3.0473 | 2.9 | 2.9 | 3.06 | 2.9 | 3 | 3 | 2.9 | 3 |  |  |  | 1
| 1 | 2 | 3.04 | 3 | 3.042 | 4 | 2.1 | 2.93 | 2.96 | 3.09 | 3.02 | 2.83 | 3.17 | 3.2 | 2.9 | 2.9499 | 3.1 | 3 | 2.88 | 2.6 | 3 | 4 | 3.4 | 3 |  |  |  | 36
| 1 | 3 | 3 | 2.9 | 2.8311 | 4 | 2.03 | 2.94 | 2.96 | 3.08 | 3.01 | 2.91 | 3.11 | 2.82 | 3.1 | 3.044 | 2.8 | 2.8 | 2.76 | 2.7 | 3 | 4 | 3.1 | 3 |  |  |  | 16
| 1 | 4 | 3.1 | 2.9 | 2.8795 | 4 | 1.87 | 2.92 | 2.96 | 3.09 | 3.01 | 2.8 | 3.04 | 3.28 | 3.1 | 3.1125 | 2.9 | 2.8 | 2.97 | 2.6 | 3 | 4 | 3.2 | 3 |  |  |  | 2.986
| 1 | 5 | 2.91 | 2.9 | 2.8234 | 4 | 2.09 | 2.98 | 3 | 3.04 | 3.02 | 2.88 | 3.03 | 3.23 | 3.1 | 2.901 | 2.9 | 2.7 | 2.91 | 2.9 | 3 | 3 | 3.2 | 3 |  |  |  | 22
| 1 | 6 | 3.04 | 2.9 | 3.0058 | 4 | 1.88 | 2.86 | 2.95 | 3.1 | 2.94 | 2.83 | 3.01 | 2.96 | 3.1 | 3.0267 | 3.1 | 2.7 | 3.18 | 2.6 | 3 | 3 | 3 | 3
| Mean |  | 3.0067 | 2.9167 | 2.9192 | 4 | 1.985 | 2.9217 | 2.9717 | 3.0617 | 3.0017 | 2.8583 | 3.0667 | 3.1233 | 3.05 | 3.0136 | 2.95 | 2.8167 | 2.96 | 2.7167 | 3 | 3.5 | 3.1333 | 3
| Median |  | 3.02 | 2.9 | 2.9065 | 4 | 1.985 | 2.925 | 2.96 | 3.085 | 3.01 | 2.855 | 3.04 | 3.215 | 3.1 | 3.0354 | 2.9 | 2.8 | 2.94 | 2.65 | 3 | 3.5 | 3.15 | 3
| Std Dev. |  | 0.068605 | 0.040825 | 0.09089 | 0 | 0.10252 | 0.040208 | 0.022286 | 0.049565 | 0.030605 | 0.04446 | 0.060882 | 0.1879 | 0.083666 | 0.075832 | 0.12247 | 0.1169 | 0.14656 | 0.1472 | 0 | 0.54772 | 0.17512 | 0
| Rel.Std.Dev. |  | 2.282% | 1.4% | 3.114% | 0% | 5.165% | 1.376% | 0.75% | 1.619% | 1.02% | 1.555% | 1.985% | 6.016% | 2.743% | 2.516% | 4.152% | 4.15% | 4.951% | 5.418% | 0% | 15.65% | 5.589% | 0%
| PDM3 |  | 0.6919% | -2.322% | -2.237% | 33.96% | -33.52% | -2.155% | -0.4802% | 2.534% | 0.5245% | -4.276% | 2.701% | 4.599% | 2.143% | 0.9226% | -1.206% | -5.671% | -0.8709% | -9.02% | 0.4687% | 17.21% | 4.934% | 0.4687%
| Z-Score (Absolute) |  | 0.41 | 0.5 | 0.47 | NA | 9.92 | 0.45 | 0.06 | 0.97 | 0.36 | 1.09 | 1.02 | 1.59 | 0.85 | 0.48 | 0.16 | 1.51 | 0.06 | 2.52 | NA | NA | 1.69 | NA
Results from laboratories 2.04, 2.23, 2.24 and 2.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 906b (Certified Value 9.99 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 10.15 | 9.96 | 9.8509 | 9.96 | 9.56 | 9.51 | 9.66 | 10.55 | 9.9 | 10.1 | 9.68 | 9.81 | 9.65 | 9.9105 | 11 | 9.84 | 6.9682 | 11 | 11.3 | 10.83 | 10.4 | 9.25 | 9.7 | 10.57 | 10.38 | 1
| 1 | 2 | 9.98 | 10 | 9.6849 | 10.77 | 9.1 | 9.48 | 9.64 | 10.15 | 10.15 | 9.85 | 9.58 | 9.33 | 9.79 | 9.7174 | 8 | 9.98 | 6.9002 | 9 | 11.2 | 10.68 | 10.6 | 9.04 | 10.1 | 10.39 | 10.37 | 37
| 1 | 3 | 9.87 | 9.84 | 9.7337 | 9.88 | 9.53 | 9.53 | 9.68 | 10.45 | 10.15 | 9.22 | 9.81 | 9.19 | 9.69 | 9.9404 | 10 | 10.06 | 6.8994 | 10 | 11.2 | 10.7 | 10.8 | 8.69 | 10.1 | 10.3 | 9.91 | 16
| 1 | 4 | 9.94 | 10.2 | 9.8615 | 10.13 | 9.32 | 9.29 | 9.48 | 10.15 | 10.1 | 9.64 | 9.59 | 9.13 | 9.76 | 9.8297 | 8 | 10.04 | 7.0039 | 9 | 11.1 | 10.63 | 10.9 | 8.83 | 10.2 | 10.69 | 10.39 | 9.993
| 1 | 5 | 9.8 | 10 | 9.9143 | 10.45 | 9.74 | 9.55 | 9.61 | 10.75 | 10.3 | 10 | 9.46 | 9.31 | 9.6 | 10.0013 | 9 | 9.97 | 7.0005 | 9 | 11.1 | 10.54 | 10.9 | 8.94 | 10.3 | 10.64 | 10.13 | 23
| 1 | 6 | 10.16 | 10.1 | 9.9504 | 10.73 | 9.09 | 9.56 | 9.38 | 10.15 | 9.78 | 10.2 | 9.72 | 9.48 | 9.78 | 9.6309 | 11 | 9.79 | 7.0515 | 11 | 10.6 | 10.87 | 10.7 | 9.12 | 10.2 | 10.2 | 10.24
| Mean |  | 9.9833 | 10.0167 | 9.8326 | 10.32 | 9.39 | 9.4867 | 9.575 | 10.3667 | 10.0633 | 9.835 | 9.64 | 9.375 | 9.7117 | 9.8384 | 9.5 | 9.9467 | 6.9706 | 9.8333 | 11.0833 | 10.7083 | 10.7167 | 8.9783 | 10.1 | 10.465 | 10.2367
| Median |  | 9.96 | 10 | 9.8562 | 10.29 | 9.425 | 9.52 | 9.625 | 10.3 | 10.125 | 9.925 | 9.635 | 9.32 | 9.725 | 9.8701 | 9.5 | 9.975 | 6.9844 | 9.5 | 11.15 | 10.69 | 10.75 | 8.99 | 10.15 | 10.48 | 10.305
| Std Dev. |  | 0.14652 | 0.12291 | 0.10327 | 0.38647 | 0.26458 | 0.10053 | 0.11895 | 0.25626 | 0.18938 | 0.35976 | 0.1228 | 0.24525 | 0.077309 | 0.14136 | 1.3784 | 0.10875 | 0.060956 | 0.98319 | 0.24833 | 0.12352 | 0.19408 | 0.20213 | 0.20976 | 0.19766 | 0.1897
| Rel.Std.Dev. |  | 1.468% | 1.227% | 1.05% | 3.745% | 2.818% | 1.06% | 1.242% | 2.472% | 1.882% | 3.658% | 1.274% | 2.616% | 0.796% | 1.437% | 14.51% | 1.093% | 0.8745% | 9.999% | 2.241% | 1.153% | 1.811% | 2.251% | 2.077% | 1.889% | 1.853%
| PDM3 |  | -0.09672% | 0.2368% | -1.605% | 3.272% | -6.034% | -5.067% | -4.183% | 3.739% | 0.7038% | -1.581% | -3.532% | -6.184% | -2.815% | -1.547% | -4.933% | -0.4636% | -30.25% | -1.598% | 10.91% | 7.158% | 7.242% | -10.15% | 1.071% | 4.723% | 2.438%
| Z-Score (Absolute) |  | 0.07 | 0.13 | 0.21 | 0.68 | 1.01 | 0.83 | 0.67 | 0.76 | 0.21 | 0.2 | 0.56 | 1.04 | 0.43 | 0.2 | NA | 0 | 5.4 | NA | 2.06 | 1.38 | 1.4 | 1.76 | 0.28 | 0.94 | 0.53
Results from laboratories 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 906b (Certified Value 0.644 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.6457 | 0.632 | 0.64699 | 0.66 | 0.63 | 0.65 | 0.65 | 0.67 | 0.64 | 0.66 | 0.628 | 0.71 | 0.6 | 0.65217 | 0.659 | 0.6 | 0.64928 | 0.61 | 0.593 | 0.61 | 0.546 | 0.56 | 0.63 | 0.6688 | 0.64 | 1
| 1 | 2 | 0.6449 | 0.636 | 0.63365 | 0.68 | 0.62 | 0.66 | 0.62 | 0.68 | 0.64 | 0.65 | 0.626 | 0.66 | 0.64 | 0.65732 | 0.668 | 0.6 | 0.64565 | 0.64 | 0.542 | 0.63 | 0.552 | 0.56 | 0.68 | 0.6733 | 0.62 | #N/A
| 1 | 3 | 0.6422 | 0.631 | 0.63399 | 0.67 | 0.61 | 0.66 | 0.64 | 0.68 | 0.64 | 0.63 | 0.63 | 0.67 | 0.61 | 0.65741 | 0.65 | 0.6 | 0.6465 | 0.62 | 0.595 | 0.62 | 0.564 | 0.56 | 0.71 | 0.6689 | 0.6 | 16
| 1 | 4 | 0.6435 | 0.635 | 0.64301 | 0.66 | 0.62 | 0.66 | 0.64 | 0.68 | 0.65 | 0.64 | 0.619 | 0.686 | 0.6 | 0.6533 | 0.658 | 0.6 | 0.64719 | 0.63 | 0.545 | 0.62 | 0.556 | 0.55 | 0.66 | 0.6652 | 0.62 | 0.64419
| 1 | 5 | 0.6427 | 0.633 | 0.64088 | 0.67 | 0.62 | 0.67 | 0.64 | 0.67 | 0.64 | 0.66 | 0.629 | 0.69 | 0.61 | 0.65127 | 0.644 | 0.6 | 0.64791 | 0.63 | 0.579 | 0.62 | 0.566 | 0.55 | 0.62 | 0.6641 | 0.61 | 24
| 1 | 6 | 0.6437 | 0.633 | 0.63236 | 0.68 | 0.61 | 0.64 | 0.63 | 0.68 | 0.63 | 0.69 | 0.629 | 0.71 | 0.61 | 0.65751 | 0.65 | 0.6 | 0.64778 | 0.63 | 0.546 | 0.57 | 0.564 | 0.55 | 0.61 | 0.6584 | 0.61
| Mean |  | 0.64378 | 0.63333 | 0.63848 | 0.67 | 0.61833 | 0.65667 | 0.63667 | 0.67667 | 0.64 | 0.655 | 0.62683 | 0.68767 | 0.61167 | 0.65483 | 0.65483 | 0.6 | 0.64738 | 0.62667 | 0.56667 | 0.61167 | 0.558 | 0.555 | 0.65167 | 0.66645 | 0.61667
| Median |  | 0.6436 | 0.633 | 0.63744 | 0.67 | 0.62 | 0.66 | 0.64 | 0.68 | 0.64 | 0.655 | 0.6285 | 0.688 | 0.61 | 0.65531 | 0.654 | 0.6 | 0.64748 | 0.63 | 0.5625 | 0.62 | 0.56 | 0.555 | 0.645 | 0.667 | 0.615
| Std Dev. |  | 0.0013182 | 0.0018619 | 0.005993 | 0.0089443 | 0.0075277 | 0.010328 | 0.010328 | 0.005164 | 0.0063246 | 0.020736 | 0.0040702 | 0.020412 | 0.01472 | 0.002902 | 0.0085421 | 0 | 0.0012519 | 0.010328 | 0.025113 | 0.02137 | 0.008 | 0.0054772 | 0.038687 | 0.0051056 | 0.013663
| Rel.Std.Dev. |  | 0.2048% | 0.294% | 0.9386% | 1.335% | 1.217% | 1.573% | 1.622% | 0.7631% | 0.9882% | 3.166% | 0.6493% | 2.968% | 2.406% | 0.4432% | 1.304% | 0% | 0.1934% | 1.648% | 4.432% | 3.494% | 1.434% | 0.9869% | 5.937% | 0.7661% | 2.216%
| PDM3 |  | -0.06266% | -1.685% | -0.8862% | 4.007% | -4.013% | 1.937% | -1.167% | 5.042% | -0.65% | 1.679% | -2.694% | 6.75% | -5.048% | 1.652% | 1.653% | -6.859% | 0.4964% | -2.72% | -12.03% | -5.048% | -13.38% | -13.84% | 1.161% | 3.456% | -4.272%
| Z-Score (Absolute) |  | 0.18 | 0.24 | 0.03 | 1.25 | 0.85 | 0.71 | 0.1 | 1.52 | 0.03 | 0.64 | 0.5 | 1.97 | 1.12 | 0.63 | 0.63 | NA | 0.33 | 0.51 | 2.95 | 1.12 | 3.3 | 3.42 | 0.5 | 1.11 | 0.92
Results from laboratory 2.17 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Cd in OREAS 906b (Certified Value 0.42 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.4 | 0.38 | 0.42823 | 0.4 | 0.43 | 0.41 | 0.38 | 0.39 | 0.42 | 0.38 | 0.42 | 0.53 | 0.44 | 0.40348 | < 1 | 0.69 | < 1 | 0.35 | 0.7 | 0.4 | 0.44 | 0.5 | 0.47 | 0.38 |  | 1
| 1 | 2 | 0.44 | 0.37 | 0.42555 | 0.51 | 0.54 | 0.42 | 0.39 | 0.38 | 0.39 | 0.35 | 0.45 | 0.48 | 0.43 | 0.40011 | < 1 | 0.69 | < 1 | 0.39 | 0.76 | 0.42 | 0.43 | 0.4 | 0.43 | 0.39 |  | 12
| 1 | 3 | 0.37 | 0.41 | 0.42424 | 0.43 | 0.56 | 0.4 | 0.4 | 0.41 | 0.46 | 0.46 | 0.42 | 0.47 | 0.45 | 0.39118 | < 1 | 0.64 | < 1 | 0.34 | 0.73 | 0.42 | 0.4 | < 0.3 | 0.47 | 0.37 |  | 16
| 1 | 4 | 0.37 | 0.41 | 0.46577 | 0.44 | 0.43 | 0.44 | 0.39 | 0.4 | 0.44 | 0.37 | 0.4 | 0.49 | 0.44 | 0.3904 | < 1 | 0.71 | < 1 | 0.4 | 0.68 | 0.4 | 0.43 | 0.3 | 0.44 | 0.38 |  | 0.41753
| 1 | 5 | 0.41 | 0.36 | 0.44733 | 0.46 | 0.43 | 0.42 | 0.41 | 0.41 | 0.45 | 0.42 | 0.4 | 0.45 | 0.45 | 0.36845 | < 1 | 0.71 | < 1 | 0.37 | 0.63 | 0.41 | 0.41 | 0.5 | 0.44 | 0.37 |  | 25
| 1 | 6 | 0.45 | 0.41 | 0.47089 | 0.52 | 0.61 | 0.44 | 0.38 | 0.36 | 0.42 | 0.39 | 0.42 | 0.49 | 0.45 | 0.35793 | < 1 | 0.63 | < 1 | 0.32 | 0.7 | 0.42 | 0.42 | 0.4 | 0.41 | 0.36
| Mean |  | 0.40667 | 0.39 | 0.44367 | 0.46 | 0.5 | 0.42167 | 0.39167 | 0.39167 | 0.43 | 0.395 | 0.41833 | 0.485 | 0.44333 | 0.38526 |  | 0.67833 |  | 0.36167 | 0.7 | 0.41167 | 0.42167 | 0.42 | 0.44333 | 0.375
| Median |  | 0.405 | 0.395 | 0.43778 | 0.45 | 0.485 | 0.42 | 0.39 | 0.395 | 0.43 | 0.385 | 0.42 | 0.485 | 0.445 | 0.39079 |  | 0.69 |  | 0.36 | 0.7 | 0.415 | 0.425 | 0.4 | 0.44 | 0.375
| Std Dev. |  | 0.033862 | 0.022804 | 0.020914 | 0.046904 | 0.08 | 0.016021 | 0.01169 | 0.019408 | 0.025298 | 0.03937 | 0.018348 | 0.026646 | 0.008165 | 0.018132 |  | 0.034881 |  | 0.030605 | 0.044272 | 0.0098319 | 0.01472 | 0.083666 | 0.023381 | 0.010488
| Rel.Std.Dev. |  | 8.327% | 5.847% | 4.714% | 10.2% | 16% | 3.799% | 2.985% | 4.955% | 5.883% | 9.967% | 4.386% | 5.494% | 1.842% | 4.706% |  | 5.142% |  | 8.462% | 6.325% | 2.388% | 3.491% | 19.92% | 5.274% | 2.797%
| PDM3 |  | -2.602% | -6.594% | 6.259% | 10.17% | 19.75% | 0.9905% | -6.195% | -6.195% | 2.986% | -5.396% | 0.1921% | 16.16% | 6.18% | -7.729% |  | 62.46% |  | -13.38% | 67.65% | -1.405% | 0.9905% | 0.5913% | 6.18% | -10.19%
| Z-Score (Absolute) |  | 0.34 | 0.71 | 0.49 | 0.86 | 1.76 | 0 | 0.67 | 0.67 | 0.19 | 0.6 | 0.07 | 1.42 | 0.49 | 0.82 | 1.76 | 5.77 | 1.76 | 1.35 | 6.26 | 0.22 | 0 | NA | 0.49 | 1.05
Results from laboratory 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 906b (Certified Value 92 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.19 | 2.20 | 2.23 | 2.24 | 2.25 | 2.27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 0
1
| 1 | 1 | 92.97 | 95 | 84.67 | 100 | 88.8 | 96.5 | 102 | 89.4 | 91.99 | 86.5 | 88.53 | 94.8614 | 66.4786 | 88 | 90.4 | 95.73 | 53.3 | 81.5 | 94.15 | [6 empty] | 1
| 1 | 2 | 92.26 | 96.5 | 96.61 | 100 | 84.9 | 90.1 | 101.5 | 86.2 | 92.57 | 59 | 89.94 | 95.9587 | 66.4906 | 89 | 86.9 | 96.73 | 54.2 | 91.4 | 94.51 | [6 empty] | 39
| 1 | 3 | 92.05 | 96.6 | 97.49 | 102 | 87.5 | 91.8 | 102 | 91.1 | 90.56 | 80.7 | 87.09 | 98.8429 | 66.5322 | 89 | 87.9 | 97.67 | 57.1 | 92.7 | 90.15 | [6 empty] | 16
| 1 | 4 | 91.62 | 98.3 | 87.67 | 100.5 | 85.4 | 92.2 | 102.5 | 85.1 | 93.6 | 80.4 | 90.76 | 97.3602 | 66.416 | 90 | 89.2 | 98 | 58.2 | 91.2 | 94.24 | [6 empty] | 92.1932
| 1 | 5 | 94.06 | 96.9 | 97.6 | 101 | 85.8 | 94.7 | 104 | 88.4 | 90.79 | 84 | 86.98 | 96.7181 | 68.4516 | 90 | 89.1 | 97.7 | 51.3 | 80.8 | 94.17 | [6 empty] | 26
| 1 | 6 | 93.94 | 95.4 | 94.08 | 100.5 | 84.4 | 91.9 | 97 | 87.5 | 90.7 | 82.5 | 87.71 | 98.6232 | 66.5028 | 90 | 87.8 | 97 | 62.8 | 84.6 | 89.83
| Mean |  | 92.8167 | 96.45 | 93.02 | 100.6667 | 86.1333 | 92.8667 | 101.5 | 87.95 | 91.7017 | 78.85 | 88.5017 | 97.0607 | 66.812 | 89.3333 | 88.55 | 97.1383 | 56.15 | 87.0333 | 92.8417
| Median |  | 92.615 | 96.55 | 95.345 | 100.5 | 85.6 | 92.05 | 102 | 87.95 | 91.39 | 81.6 | 88.12 | 97.0391 | 66.4967 | 89.5 | 88.5 | 97.335 | 55.65 | 87.9 | 94.16
| Std Dev. |  | 1.016 | 1.1709 | 5.5374 | 0.75277 | 1.6825 | 2.3123 | 2.3664 | 2.1732 | 1.2312 | 9.9821 | 1.5559 | 1.5412 | 0.80417 | 0.8165 | 1.2534 | 0.83784 | 4.1176 | 5.3653 | 2.215
| Rel.Std.Dev. |  | 1.095% | 1.214% | 5.953% | 0.7478% | 1.953% | 2.49% | 2.331% | 2.471% | 1.343% | 12.66% | 1.758% | 1.588% | 1.204% | 0.914% | 1.415% | 0.8625% | 7.333% | 6.165% | 2.386%
| PDM3 |  | 0.6763% | 4.617% | 0.8968% | 9.191% | -6.573% | 0.7305% | 10.09% | -4.602% | -0.5331% | -14.47% | -4.004% | 5.28% | -27.53% | -3.102% | -3.952% | 5.364% | -39.1% | -5.597% | 0.7034%
| Z-Score (Absolute) |  | 0.16 | 0.69 | 0.19 | 1.29 | 0.8 | 0.17 | 1.42 | 0.54 | 0 | 1.86 | 0.46 | 0.77 | 3.6 | 0.34 | 0.46 | 0.79 | 5.14 | 0.67 | 0.16
| Analytical results for Co in OREAS 906b (Certified Value 17.6 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 18.2 | 17 | 17.6192 | 17.3 | 15.32 | 18.5 | 17.7 | 18.2 | 18.3 | 16.6 | 17.9 | 18.4 | 17.8 | 17.7647 | 17 | 16.97 | 13.1065 | 17 | 16.2 | 16.7 | 23.6 | 18.4 | 17 | 17.1 | 18.8 | 1
| 1 | 2 | 18.1 | 17 | 17.4914 | 17.4 | 16.88 | 18.8 | 17.4 | 18.1 | 18.9 | 15.7 | 17.9 | 18.1 | 18.1 | 17.9508 | 17 | 17.52 | 12.8549 | 17 | 16.2 | 17.1 | 23.9 | 18.2 | 17 | 16.7 | 18.4 | 27
| 1 | 3 | 17.9 | 17 | 17.3369 | 17.6 | 17.87 | 18.7 | 17.6 | 18.9 | 19.1 | 16.6 | 17.9 | 17.7 | 17.6 | 17.8247 | 17 | 17.19 | 13.0019 | 17 | 16.3 | 17 | 22.9 | 17.4 | 16 | 16.6 | 18.4 | 16
| 1 | 4 | 18 | 17 | 17.2074 | 17.2 | 15.54 | 21.5 | 17.3 | 18.3 | 18.8 | 15.5 | 17.7 | 17.6 | 18.2 | 17.6497 | 17 | 17.75 | 13.1103 | 18 | 16.5 | 17.2 | 24.1 | 18 | 15 | 17.4 | 17.9 | 17.5817
| 1 | 5 | 17.8 | 17 | 17.8134 | 17.8 | 15.08 | 19 | 17.6 | 18.3 | 19 | 16.2 | 17.2 | 18.1 | 18.1 | 17.598 | 16 | 17.39 | 14.6975 | 18 | 16.9 | 16.9 | 23.6 | 17.8 | 17 | 17.3 | 17.7 | 27
| 1 | 6 | 18.4 | 17 | 17.7523 | 17.4 | 17.68 | 18.6 | 17.4 | 18.1 | 18.3 | 16.1 | 17.3 | 18 | 18 | 17.6963 | 17 | 17.32 | 13.2178 | 17 | 15.8 | 17.2 | 23.8 | 18 | 17 | 16.3 | 17.8
| Mean |  | 18.0667 | 17 | 17.5367 | 17.45 | 16.395 | 19.1833 | 17.5 | 18.3167 | 18.7333 | 16.1167 | 17.65 | 17.9833 | 17.9667 | 17.7473 | 16.8333 | 17.3567 | 13.3315 | 17.3333 | 16.3167 | 17.0167 | 23.65 | 17.9667 | 16.5 | 16.9 | 18.1667
| Median |  | 18.05 | 17 | 17.5553 | 17.4 | 16.21 | 18.75 | 17.5 | 18.25 | 18.85 | 16.15 | 17.8 | 18.05 | 18.05 | 17.7305 | 17 | 17.355 | 13.1084 | 17 | 16.25 | 17.05 | 23.7 | 18 | 17 | 16.9 | 18.15
| Std Dev. |  | 0.21602 | 0 | 0.23675 | 0.21679 | 1.2392 | 1.1479 | 0.15492 | 0.29944 | 0.35024 | 0.4535 | 0.32094 | 0.29269 | 0.22509 | 0.12815 | 0.40825 | 0.26875 | 0.68031 | 0.5164 | 0.3656 | 0.19408 | 0.41352 | 0.34448 | 0.83666 | 0.43359 | 0.43205
| Rel.Std.Dev. |  | 1.196% | 0% | 1.35% | 1.242% | 7.558% | 5.984% | 0.8853% | 1.635% | 1.87% | 2.814% | 1.818% | 1.628% | 1.253% | 0.7221% | 2.425% | 1.548% | 5.103% | 2.979% | 2.241% | 1.141% | 1.749% | 1.917% | 5.071% | 2.566% | 2.378%
| PDM3 |  | 2.758% | -3.308% | -0.2556% | -0.749% | -6.75% | 9.11% | -0.4646% | 4.18% | 6.55% | -8.333% | 0.3885% | 2.284% | 2.19% | 0.9422% | -4.256% | -1.28% | -24.17% | -1.413% | -7.195% | -3.214% | 34.51% | 2.19% | -6.152% | -3.877% | 3.327%
| Z-Score (Absolute) |  | 0.54 | NA | 0.15 | 0.26 | 1.64 | 2 | 0.2 | 0.87 | 1.41 | 2 | 0 | 0.44 | 0.41 | 0.13 | NA | 0.38 | 5.64 | NA | 1.74 | 0.83 | 7.83 | 0.41 | NA | 0.98 | 0.67
Results from laboratories 2.02, 2.16, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 906b (Certified Value 10.2 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 11 | 10 | 11.3226 | 8 | 20.25 | 13 | 11 | 12 | 10 | 9 | 9 | 9 | 10 | 10.2948 | 10 | 10.3 | 9 | 9 | 10 | NA | 10 | 18 | 10 |  |  | 1
| 1 | 2 | 13 | 9 | 11.6422 | 8 | 23.28 | 11 | 10 | 14 | 10 | 10 | 9 | 9 | 10 | 10.151 | 10 | 10.1 | 8 | 6 | 10 | NA | 10 | 14 | 10 |  |  | 13
| 1 | 3 | 11 | 11 | 12.4628 | 10 | 26.27 | 11 | 10 | 13 | 11 | 9 | 8 | 11 | 10 | 10.1854 | 10 | 10.5 | 8 | 8 | 10 | NA | 10 | 20 | 10 |  |  | 16
| 1 | 4 | 11 | 10 | 10.7783 | 8 | 27.63 | 11 | 10 | 12 | 9 | 10 | 8 | 9 | 10 | 10.1743 | 11 | 9.9 | 7 | 7 | 10 | NA | 11 | 16 | 9 |  |  | 10.1781
| 1 | 5 | 10 | 11 | 12.8542 | 10 | 24.86 | 11 | 10 | 9 | 11 | 10 | 9 | 10 | 11 | 10.3925 | 10 | 10.7 | 8 | 10 | 11 | NA | 10 | 17 | 10 |  |  | 28
| 1 | 6 | 11 | 11 | 11.1111 | 9 | 24.41 | 10 | 10 | 13 | 10 | 11 | 10 | 9 | 10 | 10.7191 | 12 | 10.4 | 10 | 7 | 10 | NA | 11 | 24 | 10
| Mean |  | 11.1667 | 10.3333 | 11.6952 | 8.8333 | 24.45 | 11.1667 | 10.1667 | 12.1667 | 10.1667 | 9.8333 | 8.8333 | 9.5 | 10.1667 | 10.3195 | 10.5 | 10.3167 | 8.3333 | 7.8333 | 10.1667 |  | 10.3333 | 18.1667 | 9.8333
| Median |  | 11 | 10.5 | 11.4824 | 8.5 | 24.635 | 11 | 10 | 12.5 | 10 | 10 | 9 | 9 | 10 | 10.2401 | 10 | 10.35 | 8 | 7.5 | 10 |  | 10 | 17.5 | 10
| Std Dev. |  | 0.98319 | 0.8165 | 0.80696 | 0.98319 | 2.5536 | 0.98319 | 0.40825 | 1.7224 | 0.75277 | 0.75277 | 0.75277 | 0.83666 | 0.40825 | 0.21589 | 0.83666 | 0.28577 | 1.0328 | 1.472 | 0.40825 |  | 0.5164 | 3.4881 | 0.40825
| Rel.Std.Dev. |  | 8.805% | 7.902% | 6.9% | 11.13% | 10.44% | 8.805% | 4.016% | 14.16% | 7.404% | 7.655% | 8.522% | 8.807% | 4.016% | 2.092% | 7.968% | 2.77% | 12.39% | 18.79% | 4.016% |  | 4.997% | 19.2% | 4.152%
| PDM3 |  | 9.712% | 1.525% | 14.91% | -13.21% | 140.2% | 9.712% | -0.1128% | 19.54% | -0.1128% | -3.388% | -13.21% | -6.663% | -0.1128% | 1.389% | 3.162% | 1.361% | -18.13% | -23.04% | -0.1128% |  | 1.525% | 78.49% | -3.388%
| Z-Score (Absolute) |  | 1.08 | 0.11 | 1.7 | 1.65 | 16.66 | 1.08 | 0.09 | 2.26 | 0.09 | 0.48 | 1.65 | 0.87 | 0.09 | 0.09 | 0.3 | 0.09 | 2.24 | 2.82 | 0.09 | NA | 0.11 | 9.29 | 0.48
| Analytical results for Cs in OREAS 906b (Certified Value 6.97 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.24 | 2.25 | 2.27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 7.17 | 6.8 | 6.8048 | 7.3 | 6.97 | 6.85 | 7.43 | 6.83 | 7.33 | 6.95 | 7.24 | 6.78 | 7.0925 | 6.7 | 4.6591 | 6.9 | 7 | 7.19 | 7.3 | [6 empty] | 1
| 1 | 2 | 7.1 | 6.9 | 6.9503 | 7.7 | 6.86 | 6.68 | 6.96 | 7.09 | 6.76 | 7.02 | 7.05 | 6.72 | 7.0162 | 6.93 | 4.6259 | 6.8 | 7 | 7.12 | 7.2 | [6 empty] | 42
| 1 | 3 | 6.98 | 6.6 | 6.7461 | 7.1 | 7.06 | 6.92 | 7.06 | 7.11 | 6.72 | 6.88 | 6.91 | 6.7 | 7.0843 | 6.75 | 4.6553 | 6.8 | 7 | 7.12 | 6.9 | [6 empty] | 16
| 1 | 4 | 7.15 | 6.9 | 6.8392 | 7.5 | 6.92 | 6.7 | 7.03 | 7.1 | 6.78 | 6.87 | 6.82 | 6.84 | 7.0906 | 6.89 | 4.6359 | 6.7 | 7 | 7.38 | 7.3 | [6 empty] | 6.97
| 1 | 5 | 7.16 | 6.8 | 7.0202 | 7.6 | 7.05 | 6.89 | 7.34 | 7 | 6.95 | 6.74 | 7.03 | 6.61 | 6.9763 | 6.93 | 4.6505 | 6.8 | 7 | 7.14 | 7.2 | [6 empty] | 29
| 1 | 6 | 7.06 | 6.7 | 7.0278 | 7.5 | 7 | 6.68 | 6.7 | 6.78 | 6.83 | 6.92 | 7.05 | 6.88 | 7.1021 | 6.81 | 4.6372 | 7 | 7 | 7.01 | 7
| Mean |  | 7.1033 | 6.7833 | 6.8981 | 7.45 | 6.9767 | 6.7867 | 7.0867 | 6.985 | 6.895 | 6.8967 | 7.0167 | 6.755 | 7.0603 | 6.835 | 4.644 | 6.8333 | 7 | 7.16 | 7.15
| Median |  | 7.125 | 6.8 | 6.8947 | 7.5 | 6.985 | 6.775 | 7.045 | 7.045 | 6.805 | 6.9 | 7.04 | 6.75 | 7.0875 | 6.85 | 4.6439 | 6.8 | 7 | 7.13 | 7.2
| Std Dev. |  | 0.073394 | 0.1169 | 0.11809 | 0.21679 | 0.077115 | 0.11201 | 0.26515 | 0.14571 | 0.2274 | 0.09395 | 0.14306 | 0.098742 | 0.051539 | 0.097108 | 0.012934 | 0.10328 | 0 | 0.1228 | 0.16432
| Rel.Std.Dev. |  | 1.033% | 1.723% | 1.712% | 2.91% | 1.105% | 1.65% | 3.742% | 2.086% | 3.298% | 1.362% | 2.039% | 1.462% | 0.73% | 1.421% | 0.2785% | 1.511% | 0% | 1.715% | 2.298%
| PDM3 |  | 1.912% | -2.679% | -1.033% | 6.886% | 0.09502% | -2.631% | 1.673% | 0.2146% | -1.077% | -1.053% | 0.6689% | -3.085% | 1.296% | -1.937% | -33.37% | -1.961% | 0.4298% | 2.725% | 2.582%
| Z-Score (Absolute) |  | 0.82 | 0.76 | 0.19 | 2.54 | 0.19 | 0.75 | 0.74 | 0.24 | 0.21 | 0.2 | 0.39 | 0.9 | 0.61 | 0.51 | 11.36 | 0.52 | NA | 1.1 | 1.05
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 906b (Certified Value 0.314 wt.%) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.31217 | 0.305 | 0.3103 | 0.34654 | 0.3434 | 0.299 | 0.313 | 0.315 | 0.317 | 0.304 | 0.3167 | 0.334 | 0.3198 | 0.31403 | 0.3075 | 0.3123 | 0.32066 | 0.3048 | 0.3232 | 0.3238 | 0.3144 | 0.3051 | 0.318 | 0.323 | 0.32797 | 1
| 1 | 2 | 0.30988 | 0.304 | 0.30438 | 0.3459 | 0.3473 | 0.304 | 0.312 | 0.318 | 0.322 | 0.299 | 0.3154 | 0.332 | 0.3198 | 0.32135 | 0.3111 | 0.3097 | 0.32157 | 0.3143 | 0.3195 | 0.3317 | 0.3192 | 0.3011 | 0.314 | 0.32788 | 0.32115 | 29
| 1 | 3 | 0.30967 | 0.306 | 0.3039 | 0.34178 | 0.3446 | 0.305 | 0.309 | 0.322 | 0.32 | 0.307 | 0.3162 | 0.319 | 0.3154 | 0.32244 | 0.303 | 0.3077 | 0.32849 | 0.31 | 0.3227 | 0.3275 | 0.3192 | 0.3051 | 0.315 | 0.32516 | 0.31054 | 16
| 1 | 4 | 0.31316 | 0.307 | 0.30897 | 0.3413 | 0.3399 | 0.299 | 0.312 | 0.321 | 0.326 | 0.295 | 0.3122 | 0.316 | 0.3239 | 0.3133 | 0.3058 | 0.3103 | 0.32185 | 0.3137 | 0.3221 | 0.3269 | 0.31385 | 0.3086 | 0.306 | 0.32303 | 0.31463 | 0.31434
| 1 | 5 | 0.30882 | 0.302 | 0.31242 | 0.34347 | 0.3458 | 0.307 | 0.315 | 0.324 | 0.323 | 0.306 | 0.3158 | 0.328 | 0.3233 | 0.31271 | 0.3011 | 0.3146 | 0.3258 | 0.3138 | 0.3228 | 0.3315 | 0.314 | 0.3094 | 0.31 | 0.32322 | 0.31148 | 30
| 1 | 6 | 0.31614 | 0.306 | 0.31084 | 0.34038 | 0.3469 | 0.293 | 0.314 | 0.319 | 0.313 | 0.297 | 0.3151 | 0.328 | 0.3156 | 0.31236 | 0.3017 | 0.3059 | 0.32383 | 0.3117 | 0.3161 | 0.3243 | 0.3128 | 0.3061 | 0.311 | 0.32149 | 0.31057
| Mean |  | 0.31164 | 0.305 | 0.30847 | 0.34323 | 0.34465 | 0.30117 | 0.3125 | 0.31983 | 0.32017 | 0.30133 | 0.31523 | 0.32617 | 0.31963 | 0.31603 | 0.30503 | 0.31008 | 0.3237 | 0.31138 | 0.32107 | 0.32762 | 0.31557 | 0.3059 | 0.31233 | 0.32396 | 0.31606
| Median |  | 0.31102 | 0.3055 | 0.30964 | 0.34263 | 0.3452 | 0.3015 | 0.3125 | 0.32 | 0.321 | 0.3015 | 0.3156 | 0.328 | 0.3198 | 0.31366 | 0.3044 | 0.31 | 0.32284 | 0.3127 | 0.3224 | 0.3272 | 0.3142 | 0.3056 | 0.3125 | 0.32312 | 0.31305
| Std Dev. |  | 0.0027485 | 0.0017889 | 0.0035358 | 0.0025335 | 0.0027399 | 0.0051543 | 0.0020736 | 0.0031885 | 0.0046224 | 0.0050067 | 0.0015908 | 0.0071671 | 0.0036302 | 0.0045874 | 0.0038521 | 0.0031205 | 0.0029839 | 0.0036074 | 0.0027718 | 0.0034014 | 0.0028572 | 0.0029631 | 0.0042269 | 0.0022468 | 0.0070907
| Rel.Std.Dev. |  | 0.882% | 0.5865% | 1.146% | 0.7381% | 0.795% | 1.711% | 0.6636% | 0.9969% | 1.444% | 1.662% | 0.5046% | 2.197% | 1.136% | 1.452% | 1.263% | 1.006% | 0.9218% | 1.159% | 0.8633% | 1.038% | 0.9054% | 0.9687% | 1.353% | 0.6935% | 2.243%
| PDM3 |  | -0.8598% | -2.972% | -1.869% | 9.189% | 9.641% | -4.192% | -0.5862% | 1.747% | 1.853% | -4.139% | 0.2833% | 3.761% | 1.683% | 0.5372% | -2.962% | -1.355% | 2.977% | -0.9415% | 2.139% | 4.223% | 0.392% | -2.686% | -0.6393% | 3.061% | 0.5452%
| Z-Score (Absolute) |  | 0.48 | 1.3 | 0.87 | 3.4 | 3.57 | 1.77 | 0.38 | 0.52 | 0.56 | 1.75 | 0.04 | 1.3 | 0.5 | 0.06 | 1.29 | 0.67 | 1 | 0.51 | 0.67 | 1.48 | 0 | 1.19 | 0.4 | 1.03 | 0.06
| Analytical results for Dy in OREAS 906b (Certified Value 3.71 ppm) | [26 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
3
| 1 | 1 | 3.63 | 4 | 3.6477 | 2.6423 | 3.78 | 3.2 | 3.9 | [18 empty] | 1
| 1 | 2 | 3.57 | 4.1 | 3.621 | 2.7567 | 3.58 | 3.8 | 3.9 | [18 empty] | 16
| 1 | 3 | 3.54 | 3.9 | 3.6908 | 2.7165 | 3.39 | 3.7 | 3.8 | [18 empty] | 16
| 1 | 4 | 3.56 | 4.1 | 3.6365 | 2.6846 | 3.63 | 3.5 | 3.8 | [18 empty] | 3.705
| 1 | 5 | 3.53 | 4.2 | 3.6298 | 2.6567 | 3.69 | 3.2 | 3.9 | [18 empty] | 31
| 1 | 6 | 3.61 | 4.2 | 3.725 | 2.735 | 3.72 | 3.2 | 3.8
| Mean |  | 3.5733 | 4.0833 | 3.6585 | 2.6986 | 3.6317 | 3.4333 | 3.85
| Median |  | 3.565 | 4.1 | 3.6421 | 2.7005 | 3.66 | 3.35 | 3.85
| Std Dev. |  | 0.039328 | 0.1169 | 0.040754 | 0.045046 | 0.13732 | 0.27325 | 0.054772
| Rel.Std.Dev. |  | 1.101% | 2.863% | 1.114% | 1.669% | 3.781% | 7.959% | 1.423%
| PDM3 |  | -3.554% | 10.21% | -1.257% | -27.16% | -1.98% | -7.333% | 3.913%
| Z-Score (Absolute) |  | 0.2 | 1.54 | 0.09 | 3.17 | 0 | 0.67 | 0.74
| Analytical results for Er in OREAS 906b (Certified Value 1.07 ppm) | [26 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
3
| 1 | 1 | 1.06 | 1.1 | 1.0542 | 0.85361 | 1.06 | 0.9 | 1.1 | [18 empty] | 1
| 1 | 2 | 1.04 | 1.3 | 1.0908 | 0.8576 | 1.03 | 1.1 | 1.1 | [18 empty] | 17
| 1 | 3 | 1.05 | 1.1 | 1.082 | 0.83836 | 1.08 | 1 | 1 | [18 empty] | 16
| 1 | 4 | 1.04 | 1.2 | 1.1345 | 0.88885 | 1.06 | 1 | 1.1 | [18 empty] | 1.0734
| 1 | 5 | 1.08 | 1.2 | 1.0771 | 0.90968 | 1.06 | 0.9 | 1.1 | [18 empty] | 32
| 1 | 6 | 1.07 | 1.2 | 1.1176 | 0.87923 | 0.99 | 1 | 1
| Mean |  | 1.0567 | 1.1833 | 1.0927 | 0.87122 | 1.0467 | 0.98333 | 1.0667
| Median |  | 1.055 | 1.2 | 1.0864 | 0.86841 | 1.06 | 1 | 1.1
| Std Dev. |  | 0.01633 | 0.075277 | 0.029031 | 0.026189 | 0.032042 | 0.075277 | 0.05164
| Rel.Std.Dev. |  | 1.545% | 6.361% | 2.657% | 3.006% | 3.061% | 7.655% | 4.841%
| PDM3 |  | -1.563% | 10.24% | 1.792% | -18.84% | -2.495% | -8.395% | -0.6317%
| Z-Score (Absolute) |  | 0 | 2.37 | 0.67 | 3.47 | 0.19 | 1.37 | 0.19
| Analytical results for Eu in OREAS 906b (Certified Value 1.51 ppm) | [26 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
2
| 1 | 1 | 1.46 | 1.7 | 1.5368 | 1.6873 | 1.48 | 1.21 | 1.3 | [18 empty] | 1
| 1 | 2 | 1.43 | 1.9 | 1.5855 | 1.7259 | 1.39 | 1.43 | 1.3 | [18 empty] | 18
| 1 | 3 | 1.4 | 1.7 | 1.5427 | 1.7089 | 1.36 | 1.45 | 1.2 | [18 empty] | 16
| 1 | 4 | 1.41 | 1.8 | 1.5801 | 1.7246 | 1.37 | 1.4 | 1.3 | [18 empty] | 1.5052
| 1 | 5 | 1.41 | 1.8 | 1.5658 | 1.7678 | 1.42 | 1.2 | 1.3 | [18 empty] | 33
| 1 | 6 | 1.46 | 1.9 | 1.5513 | 1.7333 | 1.42 | 1.31 | 1.3
| Mean |  | 1.4283 | 1.8 | 1.5604 | 1.7246 | 1.4067 | 1.3333 | 1.2833
| Median |  | 1.42 | 1.8 | 1.5586 | 1.7253 | 1.405 | 1.355 | 1.3
| Std Dev. |  | 0.026394 | 0.089443 | 0.020014 | 0.026767 | 0.043665 | 0.11039 | 0.040825
| Rel.Std.Dev. |  | 1.848% | 4.969% | 1.283% | 1.552% | 3.104% | 8.279% | 3.181%
| PDM3 |  | -5.109% | 19.58% | 3.663% | 14.57% | -6.549% | -11.42% | -14.74%
| Z-Score (Absolute) |  | 0 | 1.9 | 0.67 | 1.51 | 0.11 | 0.49 | 0.74
| Analytical results for Fe in OREAS 906b (Certified Value 4.85 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 5.01 | 4.81 | 4.9466 | 4.94 | 4.8 | 4.81 | 4.8 | 4.88 | 4.9 | 4.86 | 4.82 | 5.28 | 4.69 | 4.8992 | 4.91 | 4.82 | 4.9318 | 4.68 | 4.83 | 4.64 | 5.46 | 4.99 | 5.09 | 4.9156 | 5.08 | 1
| 1 | 2 | 4.98 | 4.82 | 4.8761 | 4.95 | 4.8 | 4.84 | 4.68 | 4.95 | 4.93 | 4.81 | 4.81 | 5.2 | 4.71 | 4.9369 | 4.98 | 4.81 | 4.9326 | 4.83 | 4.38 | 4.86 | 5.36 | 4.97 | 5.05 | 4.9554 | 4.86 | #N/A
| 1 | 3 | 4.97 | 4.86 | 4.8556 | 4.93 | 4.79 | 4.88 | 4.72 | 4.98 | 4.92 | 4.72 | 4.83 | 5.1 | 4.9 | 4.928 | 4.82 | 4.81 | 4.9314 | 4.74 | 4.82 | 4.76 | 5.5 | 4.94 | 5.24 | 4.9843 | 4.82 | 16
| 1 | 4 | 4.95 | 4.89 | 4.933 | 4.84 | 4.8 | 4.79 | 4.72 | 4.95 | 4.93 | 4.77 | 4.77 | 5.02 | 4.7 | 4.8732 | 4.87 | 4.8 | 4.9357 | 4.79 | 4.4 | 4.55 | 5.7582 | 4.95 | 5.35 | 4.9406 | 4.92 | 4.8529
| 1 | 5 | 4.98 | 4.92 | 4.8425 | 4.89 | 4.81 | 4.92 | 4.78 | 4.94 | 4.95 | 4.9 | 4.83 | 5.21 | 4.65 | 4.8686 | 4.83 | 4.82 | 4.939 | 4.83 | 4.66 | 4.7 | 5.48 | 4.92 | 5.21 | 4.8458 | 4.85 | 34
| 1 | 6 | 4.98 | 4.84 | 4.8114 | 4.94 | 4.79 | 4.68 | 4.71 | 4.98 | 4.79 | 4.86 | 4.82 | 5.19 | 4.73 | 4.8561 | 4.83 | 4.79 | 4.9395 | 4.8 | 4.41 | 4.54 | 5.58 | 4.98 | 4.98 | 4.8661 | 4.78
| Mean |  | 4.9783 | 4.8567 | 4.8775 | 4.915 | 4.7983 | 4.82 | 4.735 | 4.9467 | 4.9033 | 4.82 | 4.8133 | 5.1667 | 4.73 | 4.8937 | 4.8733 | 4.8083 | 4.935 | 4.7783 | 4.5833 | 4.675 | 5.523 | 4.9583 | 5.1533 | 4.918 | 4.885
| Median |  | 4.98 | 4.85 | 4.8659 | 4.935 | 4.8 | 4.825 | 4.72 | 4.95 | 4.925 | 4.835 | 4.82 | 5.195 | 4.705 | 4.8862 | 4.85 | 4.81 | 4.9341 | 4.795 | 4.535 | 4.67 | 5.49 | 4.96 | 5.15 | 4.9281 | 4.855
| Std Dev. |  | 0.019408 | 0.042269 | 0.052772 | 0.042308 | 0.0075277 | 0.083187 | 0.045497 | 0.036697 | 0.05785 | 0.066633 | 0.022509 | 0.092014 | 0.087407 | 0.033266 | 0.062183 | 0.01169 | 0.0036474 | 0.058452 | 0.21342 | 0.12422 | 0.13527 | 0.026394 | 0.13721 | 0.053324 | 0.10616
| Rel.Std.Dev. |  | 0.3898% | 0.8703% | 1.082% | 0.8608% | 0.1569% | 1.726% | 0.9609% | 0.7419% | 1.18% | 1.382% | 0.4676% | 1.781% | 1.848% | 0.6798% | 1.276% | 0.2431% | 0.07391% | 1.223% | 4.656% | 2.657% | 2.449% | 0.5323% | 2.663% | 1.084% | 2.173%
| PDM3 |  | 2.584% | 0.07737% | 0.5074% | 1.279% | -1.125% | -0.6782% | -2.43% | 1.932% | 1.039% | -0.6782% | -0.8156% | 6.465% | -2.533% | 0.8398% | 0.4208% | -0.9186% | 1.691% | -1.537% | -5.555% | -3.666% | 13.81% | 2.172% | 6.191% | 1.341% | 0.6612%
| Z-Score (Absolute) |  | 0.98 | 0.2 | 0 | 0.37 | 0.77 | 0.56 | 1.39 | 0.67 | 0.25 | 0.56 | 0.63 | 2.82 | 1.44 | 0.16 | 0.04 | 0.68 | 0.56 | 0.97 | 2.87 | 1.98 | 6.3 | 0.79 | 2.69 | 0.39 | 0.07
| Analytical results for Ga in OREAS 906b (Certified Value 28.1 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 27.95 | 27.5 | 28.7544 | 28.6 | 32.55 | 30.1 | 28.3 | 29.5 | 29.3 | 27.9 | 27.38 | 36.6 | 27.3 | 28.7094 | 27.25 | 28 | 28.3 | 27 | 28.9 | 36.3 | 28 | 28.84 |  |  |  | 1
| 1 | 2 | 27.18 | 27.7 | 29.3545 | 29.8 | 32.68 | 30 | 27.9 | 29.1 | 29.3 | 26.8 | 28.42 | 38.2 | 26.9 | 28.4379 | 27.48 | 27 | 27.9 | 29.4 | 29 | 33.9 | 27 | 28.1 |  |  |  | 47
| 1 | 3 | 26.77 | 27.7 | 28.13 | 29.52 | 33.33 | 30.5 | 27.9 | 29.8 | 30 | 26.9 | 28.26 | 35 | 27 | 28.6875 | 27.52 | 28 | 27.9 | 28.6 | 28 | 38.9 | 27 | 27.03 |  |  |  | 16
| 1 | 4 | 27.23 | 27.2 | 28.2785 | 29.4 | 31.89 | 29.3 | 27.5 | 29.2 | 29.5 | 26.5 | 27.58 | 34.6 | 27.6 | 28.3898 | 27.63 | 28 | 28.2 | 27.4 | 28.3 | 38.7 | 28 | 27.58 |  |  |  | 28.1329
| 1 | 5 | 26.71 | 27.3 | 28.8769 | 29.64 | 32.41 | 30.3 | 28 | 29.2 | 29.7 | 26.9 | 26.73 | 37.8 | 26.6 | 28.2013 | 28 | 28 | 27.8 | 27.1 | 28.3 | 36.2 | 28 | 27.37 |  |  |  | 35
| 1 | 6 | 28.31 | 27.7 | 29.1302 | 29.36 | 32.16 | 30.2 | 26.5 | 29.4 | 28.9 | 27.3 | 26.7 | 37.5 | 27.3 | 27.8862 | 27.19 | 28 | 28.1 | 25.8 | 29.3 | 34.6 | 27 | 27.58
| Mean |  | 27.3583 | 27.5167 | 28.7541 | 29.3867 | 32.5033 | 30.0667 | 27.6833 | 29.3667 | 29.45 | 27.05 | 27.5117 | 36.6167 | 27.1167 | 28.3854 | 27.5117 | 27.8333 | 28.0333 | 27.55 | 28.6333 | 36.4333 | 27.5 | 27.75
| Median |  | 27.205 | 27.6 | 28.8157 | 29.46 | 32.48 | 30.15 | 27.9 | 29.3 | 29.4 | 26.9 | 27.48 | 37.05 | 27.15 | 28.4138 | 27.5 | 28 | 28 | 27.25 | 28.6 | 36.25 | 27.5 | 27.58
| Std Dev. |  | 0.6435 | 0.22286 | 0.47597 | 0.41774 | 0.49387 | 0.41312 | 0.63377 | 0.2582 | 0.37815 | 0.48888 | 0.73169 | 1.5079 | 0.35449 | 0.31051 | 0.29171 | 0.40825 | 0.19664 | 1.274 | 0.50465 | 2.053 | 0.54772 | 0.63737
| Rel.Std.Dev. |  | 2.352% | 0.8099% | 1.655% | 1.422% | 1.519% | 1.374% | 2.289% | 0.8792% | 1.284% | 1.807% | 2.66% | 4.118% | 1.307% | 1.094% | 1.06% | 1.467% | 0.7014% | 4.624% | 1.762% | 5.635% | 1.992% | 2.297%
| PDM3 |  | -2.753% | -2.19% | 2.208% | 4.457% | 15.54% | 6.874% | -1.598% | 4.386% | 4.682% | -3.849% | -2.208% | 30.16% | -3.612% | 0.8975% | -2.208% | -1.065% | -0.3538% | -2.072% | 1.779% | 29.5% | -2.25% | -1.361%
| Z-Score (Absolute) |  | 0.61 | 0.44 | 0.87 | 1.54 | 4.83 | 2.26 | 0.26 | 1.52 | 1.6 | 0.93 | 0.45 | 9.18 | 0.86 | 0.48 | 0.45 | 0.11 | 0.11 | 0.41 | 0.74 | 8.99 | 0.46 | 0.19
| Analytical results for Gd in OREAS 906b (Certified Value 5.9 ppm) | [26 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
3
| 1 | 1 | 6.37 | 6.4 | 6.0318 | 4.9879 | 6.13 | 4.8 | 6.1 | [18 empty] | 1
| 1 | 2 | 6.13 | 6.8 | 6.349 | 5.0758 | 5.84 | 6 | 6 | [18 empty] | 19
| 1 | 3 | 5.93 | 6.2 | 6.4064 | 5.0237 | 6.02 | 5.9 | 5.6 | [18 empty] | 16
| 1 | 4 | 6.06 | 6.7 | 6.3406 | 5.0013 | 5.9 | 5.5 | 6.3 | [18 empty] | 5.9026
| 1 | 5 | 5.87 | 6.8 | 6.2391 | 5.1421 | 6.08 | 4.5 | 6.1 | [18 empty] | 36
| 1 | 6 | 6.21 | 6.9 | 6.2522 | 5.1194 | 5.9 | 5.1 | 5.8
| Mean |  | 6.095 | 6.6333 | 6.2698 | 5.0583 | 5.9783 | 5.3 | 5.9833
| Median |  | 6.095 | 6.75 | 6.2964 | 5.0497 | 5.96 | 5.3 | 6.05
| Std Dev. |  | 0.18393 | 0.27325 | 0.13255 | 0.063963 | 0.11531 | 0.60332 | 0.24833
| Rel.Std.Dev. |  | 3.018% | 4.119% | 2.114% | 1.265% | 1.929% | 11.38% | 4.15%
| PDM3 |  | 3.26% | 12.38% | 6.222% | -14.3% | 1.283% | -10.21% | 1.368%
| Z-Score (Absolute) |  | 0.26 | 1.53 | 0.67 | 2.18 | 0.01 | 1.61 | 0
| Analytical results for Ge in OREAS 906b (Certified Value 0.21 ppm) | [26 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.20 | 2.25 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [11 empty] | 2
2
| 1 | 1 | 1.2 | 1.3 | 0.20748 | 0.19 | 0.23 | 0.18 | 0.12 | 0.17 | 0.19 | 0.55 | 0.5 | 0.38135 | 1.53 | 0.1 | [11 empty] | 1
| 1 | 2 | 1.4 | 1.3 | 0.22309 | 0.3 | 0.21 | 0.17 | 0.15 | 0.16 | 0.16 | 0.52 | 0.3 | 0.31606 | 1.51 | < 0.1 | [11 empty] | 49
| 1 | 3 | 1.2 | 1.3 | 0.20403 | 0.2 | 0.21 | 0.17 | 0.14 | 0.19 | 0.26 | 0.54 | 0.5 | 0.32097 | 1.51 | 0.2 | [11 empty] | 16
| 1 | 4 | 1.2 | 1.3 | 0.22026 | 0.19 | 0.21 | 0.15 | 0.15 | 0.16 | 0.19 | 0.53 | 0.3 | 0.33184 | 1.51 | 0.2 | [11 empty] | 0.20724
| 1 | 5 | 1.3 | 1.3 | 0.20353 | 0.26 | 0.23 | 0.18 | 0.13 | 0.19 | 0.17 | 0.55 | 0.5 | 0.3401 | 1.54 | 0.1 | [11 empty] | 37
| 1 | 6 | 1.3 | 1.3 | 0.20835 | 0.24 | 0.23 | 0.17 | 0.15 | 0.18 | 0.15 | 0.54 | 0.4 | 0.31656 | 1.54 | 0.2
| Mean |  | 1.2667 | 1.3 | 0.21112 | 0.23 | 0.22 | 0.17 | 0.14 | 0.175 | 0.18667 | 0.53833 | 0.41667 | 0.33448 | 1.5233 | 0.16
| Median |  | 1.25 | 1.3 | 0.20791 | 0.22 | 0.22 | 0.17 | 0.145 | 0.175 | 0.18 | 0.54 | 0.45 | 0.32641 | 1.52 | 0.2
| Std Dev. |  | 0.08165 | 0 | 0.0084338 | 0.044721 | 0.010954 | 0.010954 | 0.012649 | 0.013784 | 0.039328 | 0.01169 | 0.098319 | 0.024809 | 0.015055 | 0.054772
| Rel.Std.Dev. |  | 6.446% | 0% | 3.995% | 19.44% | 4.979% | 6.444% | 9.035% | 7.877% | 21.07% | 2.172% | 23.6% | 7.417% | 0.9883% | 34.23%
| PDM3 |  | 511.2% | 527.3% | 1.876% | 10.98% | 6.158% | -17.97% | -32.44% | -15.56% | -9.926% | 159.8% | 101.1% | 61.4% | 635.1% | -22.79%
| Z-Score (Absolute) |  | NA | NA | 0.07 | 0.23 | 0.07 | 0.71 | 1.18 | 0.64 | 0.45 | 5.05 | NA | 1.86 | 20.48 | NA
Results from laboratories 2.01, 2.02, 2.14 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 906b (Certified Value 7.15 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.23 | 2.24 | 2.25 | 2.27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 7.19 | 7.3 | 7.5997 | 7.1 | 4.93 | 7.3 | 6.9 | 7.3 | 7.8 | 6.9 | 7.01 | 7.5 | 7.57 | 6.7976 | 7.28 | NA | 7.75 | 6.7 | 7.32 | [6 empty] | 1
| 1 | 2 | 7.05 | 7.2 | 6.7232 | 7.81 | 4.89 | 7.3 | 7 | 7 | 8 | 6.5 | 6.97 | 7.1 | 7.48 | 7.1896 | 7.04 | NA | 7.81 | 2.9 | 7.15 | [6 empty] | 33
| 1 | 3 | 6.94 | 7.3 | 7.0999 | 7.32 | 4.97 | 7.3 | 7.1 | 7.2 | 7.9 | 6.7 | 6.97 | 7 | 7.67 | 7.3332 | 6.94 | NA | 7.47 | 6.4 | 6.78 | [6 empty] | 16
| 1 | 4 | 6.99 | 7.3 | 6.8477 | 7.58 | 4.87 | 7.2 | 6.9 | 7 | 7.9 | 6.5 | 6.79 | 6.95 | 7.39 | 7.3311 | 7.21 | NA | 7.72 | 6.9 | 7.11 | [6 empty] | 7.1516
| 1 | 5 | 7.06 | 7.4 | 7.6715 | 7.7 | 4.92 | 7.4 | 7 | 7.3 | 7.9 | 6.7 | 6.72 | 7.2 | 7.02 | 7.0096 | 7.12 | NA | 7.54 | 6.9 | 6.99 | [6 empty] | 38
| 1 | 6 | 7.14 | 7.4 | 7.6568 | 7.89 | 5 | 7.4 | 6.7 | 6.9 | 7.5 | 6.4 | 6.69 | 7.2 | 7.47 | 7.177 | 6.97 | NA | 7.68 | 6.9 | 6.91
| Mean |  | 7.0617 | 7.3167 | 7.2665 | 7.5667 | 4.93 | 7.3167 | 6.9333 | 7.1167 | 7.8333 | 6.6167 | 6.8583 | 7.1583 | 7.4333 | 7.1397 | 7.0933 |  | 7.6617 | 6.1167 | 7.0433
| Median |  | 7.055 | 7.3 | 7.3498 | 7.64 | 4.925 | 7.3 | 6.95 | 7.1 | 7.9 | 6.6 | 6.88 | 7.15 | 7.475 | 7.1833 | 7.08 |  | 7.7 | 6.8 | 7.05
| Std Dev. |  | 0.092394 | 0.075277 | 0.43029 | 0.30342 | 0.04858 | 0.075277 | 0.13663 | 0.17224 | 0.17512 | 0.18348 | 0.14148 | 0.196 | 0.22402 | 0.20588 | 0.13471 |  | 0.13045 | 1.588 | 0.19096
| Rel.Std.Dev. |  | 1.308% | 1.029% | 5.922% | 4.01% | 0.9854% | 1.029% | 1.971% | 2.42% | 2.236% | 2.773% | 2.063% | 2.738% | 3.014% | 2.884% | 1.899% |  | 1.703% | 25.96% | 2.711%
| PDM3 |  | -1.257% | 2.308% | 1.606% | 5.804% | -31.06% | 2.308% | -3.052% | -0.4884% | 9.533% | -7.48% | -4.101% | 0.09425% | 3.94% | -0.1664% | -0.8146% |  | 7.132% | -14.47% | -1.514%
| Z-Score (Absolute) |  | 0.23 | 0.66 | 0.49 | 1.54 | 7.73 | 0.66 | 0.69 | 0.04 | 2.48 | 1.8 | 0.95 | 0.11 | 1.07 | 0.04 | 0.12 | NA | 1.88 | 3.56 | 0.3
| Analytical results for Hg in OREAS 906b (Indicative Value < 1 ppm) | [26 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 2.15 | 2.19 | 2.20 | 2.26 | [21 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
2
| 1 | 1 | < 1 | < 2 | < 0.05 | 0.03 | [21 empty] | 1
| 1 | 2 | < 1 | < 2 | < 0.05 | 0.03 | [21 empty] | 14
| 1 | 3 | < 1 | < 2 | < 0.05 | 0.03 | [21 empty] | 16
| 1 | 4 | < 1 | < 2 | < 0.05 | 0.04 | [21 empty] | < 1
| 1 | 5 | < 1 | < 2 | < 0.05 | 0.03 | [21 empty] | 13
| 1 | 6 | < 1 | < 2 | < 0.05 | 0.02
| Mean |  |  |  |  | 0.03
| Median |  |  |  |  | 0.03
| Std Dev. |  |  |  |  | 0.0063246
| Rel.Std.Dev. |  |  |  |  | 21.08%
PDM3
| Z-Score (Absolute) |  | 0.67 | 2.09 | 0.68 | 0.67
| Analytical results for Ho in OREAS 906b (Certified Value 0.5 ppm) | [26 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
2
| 1 | 1 | 0.52 | 0.6 | 0.51287 | 0.36406 | 0.53 | 0.4 | 0.5 | [18 empty] | 1
| 1 | 2 | 0.51 | 0.6 | 0.5411 | 0.39963 | 0.51 | 0.5 | 0.5 | [18 empty] | 16
| 1 | 3 | 0.51 | 0.6 | 0.52776 | 0.37535 | 0.51 | 0.5 | 0.5 | [18 empty] | 16
| 1 | 4 | 0.5 | 0.6 | 0.54205 | 0.37968 | 0.51 | 0.5 | 0.5 | [18 empty] | 0.50291
| 1 | 5 | 0.51 | 0.6 | 0.51217 | 0.40236 | 0.53 | 0.4 | 0.5 | [18 empty] | 39
| 1 | 6 | 0.51 | 0.7 | 0.52243 | 0.38291 | 0.51 | 0.5 | 0.5
| Mean |  | 0.51 | 0.61667 | 0.5264 | 0.384 | 0.51667 | 0.46667 | 0.5
| Median |  | 0.51 | 0.6 | 0.5251 | 0.3813 | 0.51 | 0.5 | 0.5
| Std Dev. |  | 0.0063246 | 0.040825 | 0.013146 | 0.014654 | 0.010328 | 0.05164 | 0
| Rel.Std.Dev. |  | 1.24% | 6.62% | 2.497% | 3.816% | 1.999% | 11.07% | 0%
| PDM3 |  | 1.409% | 22.62% | 4.669% | -23.65% | 2.735% | -7.207% | -0.5794%
| Z-Score (Absolute) |  | 0 | 4.39 | 0.67 | 5.18 | 0.27 | 1.78 | 0.41
| Analytical results for In in OREAS 906b (Certified Value 1.16 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.18 | 1.14 | 1.1922 | 1.19 | < 1 | 1.2 | 1.16 | 1.14 | 1.29 | 1.15 | 1.167 | 1.11 | 1.07 | 1.035 | 1.07 | 1.12 | 1.16 | 1.34 | 1.22 | 1.2 | 1.18 |  |  |  |  | 1
| 1 | 2 | 1.19 | 1.14 | 1.2388 | 1.24 | < 1 | 1.2 | 1.145 | 1.13 | 1.305 | 1.04 | 1.18 | 1.02 | 1.07 | 1.0102 | 1.1 | 1.1 | 1.23 | 1.29 | 1.22 | 1.1 | 1.15 |  |  |  |  | 35
| 1 | 3 | 1.17 | 1.14 | 1.1854 | 1.15 | < 1 | 1.265 | 1.14 | 1.18 | 1.32 | 1.185 | 1.176 | 1.04 | 1.05 | 1.0569 | 1.1 | 1.09 | 1.19 | 1.29 | 1.17 | 1.2 | 1.14 |  |  |  |  | 16
| 1 | 4 | 1.17 | 1.13 | 1.181 | 1.18 | < 1 | 1.215 | 1.135 | 1.15 | 1.295 | 1.055 | 1.157 | 1.07 | 1.06 | 1.0234 | 1.1 | 1.14 | 1.19 | 1.27 | 1.2 | 1.2 | 1.17 |  |  |  |  | 1.1595
| 1 | 5 | 1.19 | 1.16 | 1.2374 | 1.24 | < 1 | 1.26 | 1.14 | 1.13 | 1.325 | 1.125 | 1.136 | 1.05 | 1.05 | 1.0087 | 1.07 | 1.19 | 1.14 | 1.2 | 1.19 | 1.2 | 1.14 |  |  |  |  | 40
| 1 | 6 | 1.18 | 1.15 | 1.212 | 1.27 | < 1 | 1.255 | 1.105 | 1.14 | 1.28 | 1.095 | 1.142 | 1.04 | 1.06 | 1.0195 | 1.07 | 1.17 | 1.21 | 1.26 | 1.21 | 1.2 | 1.15
| Mean |  | 1.18 | 1.1433 | 1.2078 | 1.2117 |  | 1.2325 | 1.1375 | 1.145 | 1.3025 | 1.1083 | 1.1597 | 1.055 | 1.06 | 1.0256 | 1.085 | 1.135 | 1.1867 | 1.275 | 1.2017 | 1.1833 | 1.155
| Median |  | 1.18 | 1.14 | 1.2021 | 1.215 |  | 1.235 | 1.14 | 1.14 | 1.3 | 1.11 | 1.162 | 1.045 | 1.06 | 1.0214 | 1.085 | 1.13 | 1.19 | 1.28 | 1.205 | 1.2 | 1.15
| Std Dev. |  | 0.0089443 | 0.010328 | 0.025771 | 0.04535 |  | 0.030781 | 0.018097 | 0.018708 | 0.017536 | 0.055827 | 0.017963 | 0.031464 | 0.0089443 | 0.018069 | 0.016432 | 0.03937 | 0.03266 | 0.045935 | 0.019408 | 0.040825 | 0.016432
| Rel.Std.Dev. |  | 0.758% | 0.9033% | 2.134% | 3.743% |  | 2.497% | 1.591% | 1.634% | 1.346% | 5.037% | 1.549% | 2.982% | 0.8438% | 1.762% | 1.514% | 3.469% | 2.752% | 3.603% | 1.615% | 3.45% | 1.423%
| PDM3 |  | 1.768% | -1.395% | 4.165% | 4.499% |  | 6.295% | -1.898% | -1.251% | 12.33% | -4.413% | 0.01402% | -9.013% | -8.582% | -11.55% | -6.426% | -2.113% | 2.343% | 9.961% | 3.636% | 2.055% | -0.3884%
| Z-Score (Absolute) |  | 0.36 | 0.17 | 0.76 | 0.82 | 9.46 | 1.12 | 0.25 | 0.14 | 2.13 | 0.67 | 0.07 | 1.44 | 1.37 | 1.87 | 1.01 | 0.29 | 0.46 | 1.73 | 0.67 | 0.41 | 0
| Analytical results for K in OREAS 906b (Certified Value 2.88 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.8814 | 2.82 | 2.8324 | 2.74 | 3.61 | 2.94 | 2.88 | 2.91 | 2.98 | 2.9 | 2.91 | 2.27 | 2.94 | 2.8783 | 2.9 | 2.84 | 2.9406 | 2.82 | 3 | 2.78 | 2.8741 | 2.94 | 3.54 | 3.4 |  | 1
| 1 | 2 | 2.8713 | 2.84 | 2.7716 | 2.7 | 3.61 | 2.96 | 2.84 | 2.98 | 3.03 | 2.89 | 2.9 | 2.34 | 2.92 | 2.9119 | 2.93 | 2.83 | 2.9424 | 2.92 | 2.73 | 2.87 | 2.9035 | 2.97 | 1.43 | 3.27 |  | #N/A
| 1 | 3 | 2.8568 | 2.88 | 2.6879 | 2.63 | 3.59 | 2.98 | 2.83 | 2.96 | 2.91 | 2.83 | 2.92 | 2.2 | 3.01 | 2.9457 | 2.87 | 2.83 | 2.9409 | 2.85 | 2.94 | 2.87 | 2.841 | 2.95 | 2.17 | 3.23 |  | 16
| 1 | 4 | 2.8439 | 2.85 | 2.8143 | 2.66 | 3.59 | 2.92 | 2.82 | 2.97 | 3.02 | 2.81 | 2.88 | 2.29 | 2.87 | 2.8857 | 2.88 | 2.82 | 2.9418 | 2.88 | 2.74 | 2.87 | 2.8956 | 2.92 | 3.49 | 3.2 |  | 2.8848
| 1 | 5 | 2.8736 | 2.9 | 2.6818 | 2.42 | 3.59 | 3 | 2.84 | 2.98 | 2.99 | 2.92 | 2.91 | 2.17 | 2.98 | 2.8556 | 2.84 | 2.81 | 2.9448 | 2.89 | 2.87 | 2.81 | 2.852 | 2.95 | 3.4 | 3.18 |  | 41
| 1 | 6 | 2.8612 | 2.83 | 2.7283 | 2.12 | 3.6 | 2.89 | 2.83 | 2.96 | 2.89 | 2.85 | 2.91 | 2.32 | 2.92 | 2.8897 | 2.85 | 2.82 | 2.948 | 2.89 | 2.75 | 2.67 | 2.8557 | 2.94 | 3.22 | 3.14
| Mean |  | 2.8647 | 2.8533 | 2.7527 | 2.545 | 3.5983 | 2.9483 | 2.84 | 2.96 | 2.97 | 2.8667 | 2.905 | 2.265 | 2.94 | 2.8945 | 2.8783 | 2.825 | 2.9431 | 2.875 | 2.8383 | 2.8117 | 2.8703 | 2.945 | 2.875 | 3.2367
| Median |  | 2.8662 | 2.845 | 2.75 | 2.645 | 3.595 | 2.95 | 2.835 | 2.965 | 2.985 | 2.87 | 2.91 | 2.28 | 2.93 | 2.8877 | 2.875 | 2.825 | 2.9421 | 2.885 | 2.81 | 2.84 | 2.8649 | 2.945 | 3.31 | 3.215
| Std Dev. |  | 0.013483 | 0.030768 | 0.063747 | 0.23611 | 0.0098319 | 0.040208 | 0.020976 | 0.026077 | 0.057619 | 0.043205 | 0.013784 | 0.067157 | 0.049396 | 0.030985 | 0.033116 | 0.010488 | 0.0028344 | 0.035071 | 0.11548 | 0.079099 | 0.025171 | 0.016432 | 0.8718 | 0.09136
| Rel.Std.Dev. |  | 0.4707% | 1.078% | 2.316% | 9.278% | 0.2732% | 1.364% | 0.7386% | 0.881% | 1.94% | 1.507% | 0.4745% | 2.965% | 1.68% | 1.07% | 1.151% | 0.3713% | 0.09631% | 1.22% | 4.069% | 2.813% | 0.8769% | 0.558% | 30.32% | 2.823%
| PDM3 |  | -0.6951% | -1.089% | -4.577% | -11.78% | 24.74% | 2.204% | -1.551% | 2.608% | 2.955% | -0.627% | 0.7019% | -21.48% | 1.915% | 0.3376% | -0.2225% | -2.071% | 2.022% | -0.3381% | -1.609% | -2.534% | -0.5007% | 2.088% | -0.3381% | 12.2%
| Z-Score (Absolute) |  | 0.12 | 0.26 | 1.46 | 3.93 | 8.61 | 0.87 | 0.42 | 1.01 | 1.13 | 0.1 | 0.36 | 7.26 | 0.77 | 0.23 | 0.04 | 0.59 | 0.81 | 0 | 0.44 | 0.75 | 0.06 | 0.83 | 0 | 4.3
| Analytical results for La in OREAS 906b (Certified Value 45 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 45.4 | 48.8 | 42.02 | 48.6 | 43.9 | 47.6 | 51 | 45.4 | 45.03 | 41.6 | 42 | 47.5175 | 45.8 | 41.1 | 32.4004 | 43 | 42.9 | 46.3 | 47.1 | 20.1 | 36.2 | 47.4 |  |  |  | 1
| 1 | 2 | 43.82 | 47.4 | 43.95 | 48 | 42.2 | 44.4 | 51 | 43.5 | 45.38 | 26.1 | 42.8 | 47.8989 | 46.8 | 42.8 | 32.4032 | 44 | 41.5 | 50 | 46.2 | 20.8 | 45.3 | 47 |  |  |  | 36
| 1 | 3 | 43.1 | 47.4 | 44.39 | 49.6 | 42.8 | 45.9 | 51.7 | 46.3 | 44.82 | 38.8 | 43.1 | 48.112 | 46.3 | 41.3 | 32.4887 | 43 | 41.5 | 48.2 | 46.4 | 24.5 | 45.4 | 43.2 |  |  |  | 16
| 1 | 4 | 43.46 | 47.4 | 43.69 | 48.7 | 42.9 | 45.9 | 51.7 | 43.3 | 44.71 | 38.8 | 42.6 | 47.6559 | 46.4 | 44 | 32.4179 | 44 | 41.8 | 47 | 46.2 | 24.5 | 41.1 | 47.8 |  |  |  | 44.9691
| 1 | 5 | 42.97 | 49.7 | 42.74 | 49.7 | 43.1 | 47.7 | 51.1 | 45 | 44.37 | 39.8 | 40.2 | 47.2985 | 45.4 | 43 | 32.4576 | 44 | 42.3 | 46.8 | 45.2 | 19.7 | 33.4 | 47.6 |  |  |  | 42
| 1 | 6 | 44.47 | 46.7 | 43.34 | 48.8 | 41.6 | 45.3 | 49 | 42.7 | 44.48 | 39.8 | 42.7 | 48.3266 | 46.6 | 42.7 | 32.4235 | 44 | 42.3 | 44.8 | 46.5 | 26.2 | 37.3 | 46.4
| Mean |  | 43.87 | 47.9 | 43.355 | 48.9 | 42.75 | 46.1333 | 50.9167 | 44.3667 | 44.7983 | 37.4833 | 42.2333 | 47.8016 | 46.2167 | 42.4833 | 32.4319 | 43.6667 | 42.05 | 47.1833 | 46.2667 | 22.6333 | 39.7833 | 46.5667
| Median |  | 43.64 | 47.4 | 43.515 | 48.75 | 42.85 | 45.9 | 51.05 | 44.25 | 44.765 | 39.3 | 42.65 | 47.7774 | 46.35 | 42.75 | 32.4207 | 44 | 42.05 | 46.9 | 46.3 | 22.65 | 39.2 | 47.2
| Std Dev. |  | 0.9246 | 1.1171 | 0.86002 | 0.64498 | 0.78677 | 1.2972 | 0.99482 | 1.4052 | 0.37016 | 5.6697 | 1.0596 | 0.38377 | 0.52313 | 1.098 | 0.034549 | 0.5164 | 0.55045 | 1.7668 | 0.6186 | 2.7595 | 4.9693 | 1.7224
| Rel.Std.Dev. |  | 2.108% | 2.332% | 1.984% | 1.319% | 1.84% | 2.812% | 1.954% | 3.167% | 0.8263% | 15.13% | 2.509% | 0.8028% | 1.132% | 2.585% | 0.1065% | 1.183% | 1.309% | 3.745% | 1.337% | 12.19% | 12.49% | 3.699%
| PDM3 |  | -2.444% | 6.518% | -3.589% | 8.741% | -4.935% | 2.589% | 13.23% | -1.34% | -0.3797% | -16.65% | -6.084% | 6.299% | 2.774% | -5.528% | -27.88% | -2.896% | -6.491% | 4.924% | 2.886% | -49.67% | -11.53% | 3.553%
| Z-Score (Absolute) |  | 0.08 | 1.2 | 0.24 | 1.52 | 0.43 | 0.64 | 2.16 | 0.08 | 0.22 | 2.11 | 0.6 | 1.17 | 0.67 | 0.52 | 3.71 | 0.14 | 0.66 | 0.97 | 0.68 | 6.82 | 1.38 | 0.78
| Analytical results for Li in OREAS 906b (Certified Value 20.6 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 20.2 | 20.1 | 21.2799 | 21.2 | 21.28 | 21 | 19.7 | 22 | 20.1 | 20.3 | 20.6 | 21 | 19 | 20.8299 | 21 | 20.7 | 24.85 | 19 | 22.3 | 17 | 22.272 | 21 | 19 | 21 |  | 1
| 1 | 2 | 20.6 | 20.3 | 21.811 | 23.1 | 20.5 | 21.1 | 19.8 | 21.9 | 20 | 20 | 20.5 | 19.9 | 20 | 21.0355 | 21 | 21.1 | 24.7 | 20 | 21.1 | 17.5 | 21.232 | 22 | 19 | 20 |  | #N/A
| 1 | 3 | 20.1 | 20.1 | 20.6126 | 21.3 | 21.18 | 21.3 | 19.7 | 22 | 19.9 | 20 | 20.2 | 19.2 | 17 | 21.4359 | 21 | 20.73 | 24.48 | 19 | 21.3 | 17.8 | 21.616 | 21 | 18 | 20 |  | 16
| 1 | 4 | 20.6 | 21.1 | 20.8221 | 22.5 | 19.05 | 21.2 | 18.9 | 22.1 | 20.1 | 19.7 | 21.2 | 19.6 | 19 | 21.2944 | 21 | 21.23 | 24.97 | 20 | 22.4 | 16 | 21.346 | 21 | 17 | 20.1 |  | 20.5939
| 1 | 5 | 20.7 | 20.1 | 21.5958 | 22.6 | 20.73 | 21.4 | 19.3 | 21.7 | 20.1 | 20.3 | 20.6 | 19.9 | 20 | 21.0573 | 20 | 20.57 | 25.17 | 20 | 21.7 | 17.5 | 21.696 | 21 | 20 | 19.8 |  | 43
| 1 | 6 | 20.4 | 21.1 | 21.5125 | 23.2 | 20.35 | 20.5 | 19.6 | 21.9 | 19.7 | 19.9 | 20.7 | 19.9 | 19 | 20.8928 | 21 | 20.2 | 24.78 | 20 | 20.7 | 15.4 | 20.832 | 21 | 20 | 19.5
| Mean |  | 20.4333 | 20.4667 | 21.2723 | 22.3167 | 20.515 | 21.0833 | 19.5 | 21.9333 | 19.9833 | 20.0333 | 20.6333 | 19.9167 | 19 | 21.091 | 20.8333 | 20.755 | 24.825 | 19.6667 | 21.5833 | 16.8667 | 21.499 | 21.1667 | 18.8333 | 20.0667
| Median |  | 20.5 | 20.2 | 21.3962 | 22.55 | 20.615 | 21.15 | 19.65 | 21.95 | 20.05 | 20 | 20.6 | 19.9 | 19 | 21.0464 | 21 | 20.715 | 24.815 | 20 | 21.5 | 17.25 | 21.481 | 21 | 19 | 20
| Std Dev. |  | 0.24221 | 0.49666 | 0.467 | 0.87044 | 0.8059 | 0.31885 | 0.34059 | 0.13663 | 0.16021 | 0.23381 | 0.3266 | 0.59805 | 1.0954 | 0.23323 | 0.40825 | 0.37163 | 0.23552 | 0.5164 | 0.67651 | 0.95847 | 0.48777 | 0.40825 | 1.169 | 0.50465
| Rel.Std.Dev. |  | 1.185% | 2.427% | 2.195% | 3.9% | 3.928% | 1.512% | 1.747% | 0.6229% | 0.8017% | 1.167% | 1.583% | 3.003% | 5.766% | 1.106% | 1.96% | 1.791% | 0.9487% | 2.626% | 3.134% | 5.683% | 2.269% | 1.929% | 6.207% | 2.515%
| PDM3 |  | -0.7796% | -0.6178% | 3.294% | 8.365% | -0.3831% | 2.377% | -5.312% | 6.504% | -2.965% | -2.722% | 0.1915% | -3.288% | -7.74% | 2.414% | 1.163% | 0.7823% | 20.55% | -4.502% | 4.805% | -18.1% | 4.395% | 2.781% | -8.549% | -2.56%
| Z-Score (Absolute) |  | 0.15 | 0.12 | 0.75 | 1.88 | 0.06 | 0.55 | 1.16 | 1.47 | 0.64 | 0.58 | 0.06 | 0.71 | 1.7 | 0.56 | 0.28 | 0.2 | 4.59 | 0.98 | 1.09 | 4 | 1 | 0.64 | 1.88 | 0.55
| Analytical results for Lu in OREAS 906b (Certified Value 0.085 ppm) | [26 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.12 | 2.14 | 2.15 | 2.20 | 2.24 | 2.25 | 2.27 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [15 empty] | 3
3
| 1 | 1 | 0.09 | < 0.1 | < 1 | 0.09 | 0.09 | 0.09003 | 0.08 | 0.08 | < 0.1 | < 0.1 | [15 empty] | 1
| 1 | 2 | 0.08 | < 0.1 | < 1 | 0.09 | 0.08 | 0.089109 | 0.08 | 0.08 | < 0.1 | < 0.1 | [15 empty] | 9
| 1 | 3 | 0.08 | < 0.1 | < 1 | 0.09 | 0.09 | 0.089446 | 0.08 | 0.08 | < 0.1 | < 0.1 | [15 empty] | 16
| 1 | 4 | 0.08 | < 0.1 | < 1 | 0.09 | 0.08 | 0.098106 | 0.08 | 0.08 | < 0.1 | < 0.1 | [15 empty] | 0.084998
| 1 | 5 | 0.08 | < 0.1 | < 1 | 0.09 | 0.08 | 0.087338 | 0.09 | 0.08 | < 0.1 | < 0.1 | [15 empty] | 44
| 1 | 6 | 0.09 | < 0.1 | < 1 | 0.09 | 0.09 | 0.085693 | 0.08 | 0.09 | < 0.1 | < 0.1
| Mean |  | 0.083333 |  |  | 0.09 | 0.085 | 0.089954 | 0.081667 | 0.081667
| Median |  | 0.08 |  |  | 0.09 | 0.085 | 0.089277 | 0.08 | 0.08
| Std Dev. |  | 0.005164 |  |  | 1.5202e-17 | 0.0054772 | 0.0042994 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 6.197% |  |  | 1.689e-14% | 6.444% | 4.78% | 4.999% | 4.999%
| PDM3 |  | -1.959% |  |  | 5.884% | 0.001994% | 5.83% | -3.92% | -3.92%
| Z-Score (Absolute) |  | 0.08 | 2.93 | 37.65 | 0.68 | 0.23 | 0.67 | 0.08 | 0.08 | 2.93 | 2.93
| Analytical results for Mg in OREAS 906b (Certified Value 0.233 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.2346 | 0.227 | 0.2344 | 0.23 | 0.24 | 0.22 | 0.22 | 0.23 | 0.22 | 0.23 | 0.234 | 0.26 | 0.22 | 0.23626 | 0.246 | 0.22 | 0.23159 | 0.22 | 0.218 | 0.18 | 0.2852 | 0.2 | 0.25 | 0.2515 | 0.26 | 1
| 1 | 2 | 0.2351 | 0.228 | 0.22916 | 0.24 | 0.24 | 0.23 | 0.21 | 0.23 | 0.22 | 0.22 | 0.235 | 0.25 | 0.21 | 0.23974 | 0.249 | 0.23 | 0.23023 | 0.23 | 0.198 | 0.2 | 0.2792 | 0.19 | 0.26 | 0.253 | 0.25 | #N/A
| 1 | 3 | 0.2358 | 0.232 | 0.23095 | 0.24 | 0.24 | 0.23 | 0.22 | 0.23 | 0.22 | 0.22 | 0.235 | 0.25 | 0.23 | 0.23856 | 0.242 | 0.22 | 0.23088 | 0.23 | 0.218 | 0.19 | 0.2804 | 0.2 | 0.27 | 0.2513 | 0.24 | 16
| 1 | 4 | 0.233 | 0.229 | 0.23377 | 0.23 | 0.24 | 0.23 | 0.22 | 0.23 | 0.22 | 0.22 | 0.231 | 0.249 | 0.23 | 0.23937 | 0.244 | 0.22 | 0.23167 | 0.23 | 0.201 | 0.2 | 0.2791 | 0.2 | 0.26 | 0.2504 | 0.24 | 0.23304
| 1 | 5 | 0.2346 | 0.239 | 0.22602 | 0.24 | 0.24 | 0.23 | 0.22 | 0.23 | 0.22 | 0.23 | 0.235 | 0.26 | 0.22 | 0.23615 | 0.241 | 0.23 | 0.23238 | 0.23 | 0.21 | 0.2 | 0.2772 | 0.2 | 0.25 | 0.2497 | 0.24 | 45
| 1 | 6 | 0.2348 | 0.224 | 0.22879 | 0.24 | 0.24 | 0.22 | 0.22 | 0.23 | 0.21 | 0.24 | 0.233 | 0.26 | 0.23 | 0.23719 | 0.243 | 0.23 | 0.2311 | 0.23 | 0.201 | 0.18 | 0.2728 | 0.2 | 0.22 | 0.2479 | 0.25
| Mean |  | 0.23465 | 0.22983 | 0.23052 | 0.23667 | 0.24 | 0.22667 | 0.21833 | 0.23 | 0.21833 | 0.22667 | 0.23383 | 0.25483 | 0.22333 | 0.23788 | 0.24417 | 0.225 | 0.23131 | 0.22833 | 0.20767 | 0.19167 | 0.27898 | 0.19833 | 0.25167 | 0.25063 | 0.24667
| Median |  | 0.2347 | 0.2285 | 0.23005 | 0.24 | 0.24 | 0.23 | 0.22 | 0.23 | 0.22 | 0.225 | 0.2345 | 0.255 | 0.225 | 0.23787 | 0.2435 | 0.225 | 0.23135 | 0.23 | 0.2055 | 0.195 | 0.27915 | 0.2 | 0.255 | 0.25085 | 0.245
| Std Dev. |  | 0.00092466 | 0.0051929 | 0.0031905 | 0.005164 | 0 | 0.005164 | 0.0040825 | 0 | 0.0040825 | 0.008165 | 0.0016021 | 0.0056716 | 0.008165 | 0.0015641 | 0.0029269 | 0.0054772 | 0.00074139 | 0.0040825 | 0.0089592 | 0.0098319 | 0.0040529 | 0.0040825 | 0.017224 | 0.0017432 | 0.008165
| Rel.Std.Dev. |  | 0.3941% | 2.259% | 1.384% | 2.182% | 0% | 2.278% | 1.87% | 0% | 1.87% | 3.602% | 0.6851% | 2.226% | 3.656% | 0.6575% | 1.199% | 2.434% | 0.3205% | 1.788% | 4.314% | 5.13% | 1.453% | 2.058% | 6.844% | 0.6955% | 3.31%
| PDM3 |  | 0.6888% | -1.378% | -1.085% | 1.554% | 2.985% | -2.737% | -6.313% | -1.307% | -6.313% | -2.737% | 0.3384% | 9.35% | -4.167% | 2.073% | 4.772% | -3.452% | -0.7451% | -2.022% | -10.89% | -17.76% | 19.71% | -14.89% | 7.991% | 7.547% | 5.845%
| Z-Score (Absolute) |  | 0.38 | 0.06 | 0 | 0.56 | 0.87 | 0.35 | 1.12 | 0.05 | 1.12 | 0.35 | 0.3 | 2.23 | 0.66 | 0.67 | 1.25 | 0.51 | 0.07 | 0.2 | 2.09 | 3.56 | 4.44 | 2.95 | 1.94 | 1.84 | 1.48
| Analytical results for Mn in OREAS 906b (Certified Value 0.032 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0322 | 0.0301 | 0.034175 | 0.0339 | 0.040205 | 0.0319 | 0.0314 | 0.0313 | 0.0334 | 0.0321 | 0.0323 | 0.035 | 0.0313 | 0.031818 | 0.0318 | 0.0314 | 0.032637 | 0.0312 | 0.0274 | 0.0303 | 0.0362 | 0.0287 | 0.0342 | 0.0336 | 0.033 | 1
| 1 | 2 | 0.0315 | 0.031 | 0.032851 | 0.0347 | 0.041018 | 0.0313 | 0.0309 | 0.0335 | 0.0338 | 0.0317 | 0.0316 | 0.0344 | 0.0313 | 0.03252 | 0.0333 | 0.0318 | 0.032624 | 0.0319 | 0.0252 | 0.0325 | 0.0378 | 0.0286 | 0.0336 | 0.0334 | 0.0317 | 38
| 1 | 3 | 0.0312 | 0.03 | 0.033199 | 0.0347 | 0.04067 | 0.0321 | 0.0316 | 0.0329 | 0.0333 | 0.0308 | 0.0318 | 0.0332 | 0.032 | 0.031655 | 0.0319 | 0.0312 | 0.032867 | 0.0317 | 0.0284 | 0.0311 | 0.0356 | 0.0283 | 0.0346 | 0.0337 | 0.0313 | 16
| 1 | 4 | 0.0322 | 0.0305 | 0.032459 | 0.0338 | 0.038821 | 0.0319 | 0.0313 | 0.0335 | 0.0336 | 0.0313 | 0.0316 | 0.0332 | 0.0319 | 0.031678 | 0.0316 | 0.0323 | 0.033465 | 0.0322 | 0.0255 | 0.0292 | 0.036 | 0.0284 | 0.0347 | 0.0335 | 0.0315 | 0.032119
| 1 | 5 | 0.0313 | 0.031 | 0.034199 | 0.0348 | 0.039609 | 0.033 | 0.0312 | 0.0322 | 0.0342 | 0.0326 | 0.0321 | 0.034 | 0.0319 | 0.031565 | 0.0324 | 0.0316 | 0.032941 | 0.0317 | 0.0267 | 0.0304 | 0.0369 | 0.0284 | 0.0351 | 0.0329 | 0.0314 | 46
| 1 | 6 | 0.0319 | 0.0299 | 0.034837 | 0.0347 | 0.040091 | 0.0304 | 0.0315 | 0.0339 | 0.0327 | 0.0335 | 0.0319 | 0.0335 | 0.0315 | 0.032056 | 0.032 | 0.0313 | 0.03236 | 0.032 | 0.0244 | 0.0303 | 0.037 | 0.0285 | 0.0328 | 0.0325 | 0.0308
| Mean |  | 0.031717 | 0.030417 | 0.03362 | 0.034433 | 0.040069 | 0.031767 | 0.031317 | 0.032883 | 0.0335 | 0.032 | 0.031883 | 0.033883 | 0.03165 | 0.031882 | 0.032167 | 0.0316 | 0.032816 | 0.031783 | 0.026267 | 0.030633 | 0.036583 | 0.028483 | 0.034167 | 0.033267 | 0.031617
| Median |  | 0.0317 | 0.0303 | 0.033687 | 0.0347 | 0.040148 | 0.0319 | 0.03135 | 0.0332 | 0.0335 | 0.0319 | 0.03185 | 0.03375 | 0.0317 | 0.031748 | 0.03195 | 0.0315 | 0.032752 | 0.0318 | 0.0261 | 0.03035 | 0.03655 | 0.02845 | 0.0344 | 0.03345 | 0.03145
| Std Dev. |  | 0.0004446 | 0.00049565 | 0.00092086 | 0.00045461 | 0.00078101 | 0.00086641 | 0.00024833 | 0.00097656 | 0.00050596 | 0.00096333 | 0.00027869 | 0.00072226 | 0.00032094 | 0.00035617 | 0.00061536 | 0.00040497 | 0.00037847 | 0.00034303 | 0.0014989 | 0.0010985 | 0.00080104 | 0.0001472 | 0.00084063 | 0.00046762 | 0.0007414
| Rel.Std.Dev. |  | 1.402% | 1.63% | 2.739% | 1.32% | 1.949% | 2.727% | 0.793% | 2.97% | 1.51% | 3.01% | 0.8741% | 2.132% | 1.014% | 1.117% | 1.913% | 1.282% | 1.153% | 1.079% | 5.706% | 3.586% | 2.19% | 0.5168% | 2.46% | 1.406% | 2.345%
| PDM3 |  | -1.251% | -5.299% | 4.675% | 7.207% | 24.75% | -1.096% | -2.497% | 2.381% | 4.301% | -0.3691% | -0.7323% | 5.495% | -1.459% | -0.7364% | 0.1499% | -1.614% | 2.171% | -1.044% | -18.22% | -4.624% | 13.9% | -11.32% | 6.377% | 3.575% | -1.563%
| Z-Score (Absolute) |  | 0.11 | 0.99 | 1.17 | 1.72 | 5.52 | 0.08 | 0.38 | 0.67 | 1.09 | 0.08 | 0 | 1.35 | 0.16 | 0 | 0.19 | 0.19 | 0.63 | 0.07 | 3.79 | 0.84 | 3.17 | 2.29 | 1.54 | 0.93 | 0.18
| Analytical results for Mo in OREAS 906b (Certified Value 4.91 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 5 | 4.8 | 4.1588 | 4.8 | 4.57 | 5.31 | 4.98 | 4.97 | 5.1 | 4.99 | 4.86 | 6.11 | 4.8 | 4.9166 | 5 | 4.79 | 5.428 | 6 | 4.2 | 5.3 | 4.2 | 5.19 | 5 | 4.7 | 5.18 | 1
| 1 | 2 | 4.9 | 4.7 | 4.6944 | 5.01 | 4.48 | 5.31 | 4.86 | 4.75 | 5.29 | 4.83 | 5.18 | 5.85 | 4.95 | 5.043 | 5 | 4.79 | 5.5995 | 6 | 4.1 | 5.7 | 4.4 | 5.27 | 6 | 4.6 | 4.78 | 39
| 1 | 3 | 5.1 | 4.7 | 4.3242 | 4.72 | 4.1 | 5.24 | 4.88 | 4.74 | 5.53 | 4.58 | 5.01 | 6.1 | 4.94 | 5.0123 | 5 | 4.99 | 5.3705 | 6 | 4.2 | 5.8 | 4.1 | 5.18 | 6 | 4.7 | 4.76 | 16
| 1 | 4 | 5 | 4.7 | 4.7198 | 5.08 | 4.33 | 5.22 | 4.78 | 4.64 | 5.27 | 4.72 | 4.91 | 5.89 | 4.97 | 4.8484 | 5 | 4.82 | 5.3405 | 6 | 4.5 | 5.3 | 4.3 | 5.22 | 5 | 5 | 4.68 | 4.9125
| 1 | 5 | 4.9 | 4.7 | 4.3544 | 5.12 | 4.29 | 5.34 | 4.83 | 4.93 | 5.64 | 5.45 | 4.8 | 5.77 | 4.88 | 4.952 | 5 | 4.96 | 5.6385 | 6 | 4.4 | 5.6 | 4.2 | 5.15 | 5 | 4.9 | 4.69 | 47
| 1 | 6 | 4.9 | 4.7 | 4.2952 | 5.07 | 4.01 | 5.19 | 4.67 | 4.59 | 5.18 | 4.93 | 4.86 | 5.76 | 4.92 | 5.1914 | 5 | 4.98 | 5.5225 | 6 | 4.2 | 5.4 | 4.1 | 5.3 | 5 | 4.6 | 4.66
| Mean |  | 4.9667 | 4.7167 | 4.4245 | 4.9667 | 4.2967 | 5.2683 | 4.8333 | 4.77 | 5.335 | 4.9167 | 4.9367 | 5.9133 | 4.91 | 4.9939 | 5 | 4.8883 | 5.4832 | 6 | 4.2667 | 5.5167 | 4.2167 | 5.2183 | 5.3333 | 4.75 | 4.7917
| Median |  | 4.95 | 4.7 | 4.3393 | 5.04 | 4.31 | 5.275 | 4.845 | 4.745 | 5.28 | 4.88 | 4.885 | 5.87 | 4.93 | 4.9821 | 5 | 4.89 | 5.4752 | 6 | 4.2 | 5.5 | 4.2 | 5.205 | 5 | 4.7 | 4.725
| Std Dev. |  | 0.08165 | 0.040825 | 0.22904 | 0.16585 | 0.21463 | 0.059805 | 0.10386 | 0.15245 | 0.20811 | 0.29984 | 0.13837 | 0.15629 | 0.061968 | 0.11886 | 0 | 0.09786 | 0.12272 | 0 | 0.15055 | 0.2137 | 0.1169 | 0.057067 | 0.5164 | 0.16432 | 0.196
| Rel.Std.Dev. |  | 1.644% | 0.8655% | 5.177% | 3.339% | 4.995% | 1.135% | 2.149% | 3.196% | 3.901% | 6.099% | 2.803% | 2.643% | 1.262% | 2.38% | 0% | 2.002% | 2.238% | 0% | 3.529% | 3.874% | 2.772% | 1.094% | 9.682% | 3.459% | 4.09%
| PDM3 |  | 1.103% | -3.986% | -9.934% | 1.103% | -12.54% | 7.243% | -1.612% | -2.901% | 8.601% | 0.08482% | 0.4919% | 20.37% | -0.05089% | 1.658% | 1.781% | -0.4919% | 11.62% | 22.14% | -13.15% | 12.3% | -14.16% | 6.226% | 8.567% | -3.308% | -2.46%
| Z-Score (Absolute) |  | 0.2 | 0.74 | 1.83 | 0.2 | 2.31 | 1.33 | 0.3 | 0.54 | 1.58 | 0.01 | 0.09 | 3.75 | 0.01 | 0.3 | NA | 0.09 | 2.13 | NA | 2.42 | 2.26 | 2.61 | 1.14 | NA | 0.61 | 0.46
Results from laboratories 2.16, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 906b (Certified Value 2.63 wt.%) | [26 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.562 | 2.52 | 2.5532 | 2.811 | 2.85 | 2.62 | 2.55 | 2.57 | 3.01 | 2.62 | 2.567 | 2.73 | 2.7 | 2.5332 | 2.389 | 2.6 | 2.727 | 2.51 | 2.62 | 2.36 | 2.548 | 2.65 | 2.83 | 2.6886 | 2.875 | 1
| 1 | 2 | 2.57 | 2.53 | 2.4953 | 2.889 | 2.81 | 2.62 | 2.55 | 2.78 | 2.6 | 2.5 | 2.567 | 2.73 | 2.73 | 2.5655 | 2.412 | 2.59 | 2.7248 | 2.6 | 2.54 | 2.43 | 2.596 | 2.67 | 2.71 | 2.7247 | 2.86 | 10
| 1 | 3 | 2.5248 | 2.54 | 2.5046 | 2.859 | 2.8 | 2.65 | 2.54 | 2.79 | 2.58 | 2.53 | 2.562 | 2.61 | 2.77 | 2.521 | 2.356 | 2.59 | 2.7233 | 2.54 | 2.55 | 2.43 | 2.52 | 2.64 | 2.94 | 2.7028 | 2.755 | 16
| 1 | 4 | 2.5334 | 2.55 | 2.5424 | 2.831 | 2.82 | 2.6 | 2.51 | 2.81 | 2.64 | 2.45 | 2.528 | 2.6 | 2.71 | 2.584 | 2.374 | 2.57 | 2.7239 | 2.56 | 2.58 | 2.41 | 2.5231 | 2.63 | 2.85 | 2.6967 | 2.735 | 2.6259
| 1 | 5 | 2.569 | 2.58 | 2.4357 | 2.863 | 2.81 | 2.67 | 2.53 | 2.64 | 2.62 | 2.55 | 2.567 | 2.7 | 2.73 | 2.5224 | 2.349 | 2.57 | 2.7245 | 2.55 | 2.62 | 2.41 | 2.544 | 2.62 | 2.88 | 2.6829 | 2.698 | 48
| 1 | 6 | 2.5387 | 2.53 | 2.4659 | 2.792 | 2.83 | 2.58 | 2.55 | 2.85 | 2.93 | 2.51 | 2.565 | 2.67 | 2.75 | 2.5595 | 2.377 | 2.57 | 2.7287 | 2.56 | 2.6 | 2.25 | 2.556 | 2.64 | 2.52 | 2.6701 | 2.697
| Mean |  | 2.5497 | 2.5417 | 2.4995 | 2.8408 | 2.82 | 2.6233 | 2.5383 | 2.74 | 2.73 | 2.5267 | 2.5593 | 2.6733 | 2.7317 | 2.5476 | 2.3762 | 2.5817 | 2.7254 | 2.5533 | 2.585 | 2.3817 | 2.5478 | 2.6417 | 2.7883 | 2.6943 | 2.77
| Median |  | 2.5503 | 2.535 | 2.5 | 2.845 | 2.815 | 2.62 | 2.545 | 2.785 | 2.63 | 2.52 | 2.566 | 2.685 | 2.73 | 2.5464 | 2.3755 | 2.58 | 2.7246 | 2.555 | 2.59 | 2.41 | 2.546 | 2.64 | 2.84 | 2.6927 | 2.745
| Std Dev. |  | 0.019711 | 0.02137 | 0.044657 | 0.036091 | 0.017889 | 0.03266 | 0.016021 | 0.10954 | 0.18868 | 0.056804 | 0.015475 | 0.057504 | 0.025626 | 0.025835 | 0.022798 | 0.013292 | 0.0020549 | 0.029439 | 0.034496 | 0.069402 | 0.027524 | 0.017224 | 0.15171 | 0.01871 | 0.078852
| Rel.Std.Dev. |  | 0.7731% | 0.8408% | 1.787% | 1.27% | 0.6343% | 1.245% | 0.6312% | 3.998% | 6.911% | 2.248% | 0.6046% | 2.151% | 0.9381% | 1.014% | 0.9595% | 0.5148% | 0.0754% | 1.153% | 1.334% | 2.914% | 1.08% | 0.652% | 5.441% | 0.6944% | 2.847%
| PDM3 |  | -2.903% | -3.207% | -4.812% | 8.185% | 7.392% | -0.09742% | -3.334% | 4.346% | 3.965% | -3.779% | -2.535% | 1.807% | 4.028% | -2.982% | -9.51% | -1.684% | 3.788% | -2.763% | -1.557% | -9.301% | -2.972% | 0.6008% | 6.186% | 2.605% | 5.488%
| Z-Score (Absolute) |  | 0.28 | 0.34 | 0.67 | 2.02 | 1.85 | 0.3 | 0.37 | 1.22 | 1.14 | 0.46 | 0.2 | 0.7 | 1.16 | 0.3 | 1.65 | 0.03 | 1.11 | 0.25 | 0 | 1.6 | 0.29 | 0.45 | 1.6 | 0.86 | 1.46
| Analytical results for Nb in OREAS 906b (Certified Value 18.3 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 19.18 | 17 | 17.6887 | 10.98 | 11.52 | 18.5 | 18.2 | 19.4 | 18.8 | 17.8 | 19.3 | 20.2 | 17.3 | 18.1476 | 17 | 19.3 | 12 | 18.144 | 16.1 | 19.4 | 18.01 |  |  |  |  | 1
| 1 | 2 | 19.69 | 17 | 17.5232 | 11.37 | 12.24 | 18.3 | 18.2 | 18.8 | 19.3 | 16.8 | 19.6 | 22.4 | 17.4 | 17.8198 | 18 | 18.8 | 11 | 18.32 | 17.1 | 13.1 | 17.25 |  |  |  |  | 40
| 1 | 3 | 19.01 | 16.8 | 17.5106 | 11.17 | 11.86 | 18.8 | 18.5 | 19.3 | 19.4 | 18.7 | 19.3 | 19.4 | 17.6 | 18.2935 | 18 | 18.5 | 12 | 18.15 | 16.5 | 18.5 | 17.27 |  |  |  |  | 16
| 1 | 4 | 18.74 | 16.7 | 17.8701 | 11.47 | 8.94 | 18.3 | 18.2 | 18.9 | 19.2 | 16.6 | 19.4 | 19.3 | 17.8 | 17.7813 | 19 | 18.6 | 12 | 16.88 | 15.9 | 20.8 | 17.31 |  |  |  |  | 18.2642
| 1 | 5 | 18.74 | 16.2 | 18.2813 | 11.56 | 10.45 | 18.9 | 18.8 | 19.4 | 19.3 | 17.3 | 19.4 | 19.7 | 17.2 | 17.739 | 18 | 19.2 | 12 | 17.52 | 16 | 20.3 | 17.66 |  |  |  |  | 49
| 1 | 6 | 19.3 | 16.5 | 18.2431 | 11.66 | 11.58 | 18.8 | 17.6 | 19.2 | 18.5 | 16.9 | 19.9 | 19.6 | 17.6 | 17.6573 | 18 | 18.8 | 12 | 16.94 | 16 | 19.6 | 17.62
| Mean |  | 19.11 | 16.7 | 17.8528 | 11.3683 | 11.0983 | 18.6 | 18.25 | 19.1667 | 19.0833 | 17.35 | 19.4833 | 20.1 | 17.4833 | 17.9064 | 18 | 18.8667 | 11.8333 | 17.659 | 16.2667 | 18.6167 | 17.52
| Median |  | 19.095 | 16.75 | 17.7794 | 11.42 | 11.55 | 18.65 | 18.2 | 19.25 | 19.25 | 17.1 | 19.4 | 19.65 | 17.5 | 17.8005 | 18 | 18.8 | 12 | 17.832 | 16.05 | 19.5 | 17.465
| Std Dev. |  | 0.36365 | 0.30984 | 0.34319 | 0.25388 | 1.2143 | 0.26833 | 0.39875 | 0.2582 | 0.35449 | 0.78677 | 0.23166 | 1.1696 | 0.22286 | 0.25347 | 0.63246 | 0.32042 | 0.40825 | 0.6413 | 0.45898 | 2.8153 | 0.29973
| Rel.Std.Dev. |  | 1.903% | 1.855% | 1.922% | 2.233% | 10.94% | 1.443% | 2.185% | 1.347% | 1.858% | 4.535% | 1.189% | 5.819% | 1.275% | 1.416% | 3.514% | 1.698% | 3.45% | 3.632% | 2.822% | 15.12% | 1.711%
| PDM3 |  | 4.631% | -8.564% | -2.252% | -37.76% | -39.23% | 1.839% | -0.0778% | 4.941% | 4.485% | -5.005% | 6.675% | 10.05% | -4.275% | -1.959% | -1.447% | 3.299% | -35.21% | -3.314% | -10.94% | 1.93% | -4.075%
| Z-Score (Absolute) |  | 0.85 | 0.85 | 0.04 | 4.59 | 4.78 | 0.49 | 0.24 | 0.88 | 0.83 | 0.39 | 1.11 | 1.54 | 0.3 | 0 | NA | 0.67 | NA | 0.17 | 1.15 | 0.5 | 0.27
Results from laboratories 2.19 and 2.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 906b (Certified Value 38.6 ppm) | [26 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.26 | 2.27 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [17 empty] | 1
2
| 1 | 1 | 40.9 | 42.7 | 37.666 | 27.8481 | 38.7 | 30.6 | 37 | 40.3 | [17 empty] | 1
| 1 | 2 | 39.82 | 45.6 | 38.9807 | 27.4315 | 36.6 | 39.1 | 36 | 40.2 | [17 empty] | 18
| 1 | 3 | 39.08 | 42.1 | 39.6781 | 27.6851 | 36.8 | 37.1 | 36 | 38.2 | [17 empty] | 16
| 1 | 4 | 39 | 44.5 | 39.2103 | 27.8254 | 37.3 | 35 | 37 | 40.3 | [17 empty] | 38.6178
| 1 | 5 | 39.11 | 44.8 | 38.77 | 27.6146 | 38.1 | 28.5 | 37 | 40.2 | [17 empty] | 50
| 1 | 6 | 40.84 | 45.5 | 38.9411 | 27.673 | 37.6 | 32.3 | 35 | 38.4
| Mean |  | 39.7917 | 44.2 | 38.8744 | 27.6796 | 37.5167 | 33.7667 | 36.3333 | 39.6
| Median |  | 39.465 | 44.65 | 38.9609 | 27.679 | 37.45 | 33.65 | 36.5 | 40.2
| Std Dev. |  | 0.88612 | 1.467 | 0.67015 | 0.15201 | 0.79352 | 4.0248 | 0.8165 | 1.01
| Rel.Std.Dev. |  | 2.227% | 3.319% | 1.724% | 0.5492% | 2.115% | 11.92% | 2.247% | 2.55%
| PDM3 |  | 3.04% | 14.46% | 0.6644% | -28.32% | -2.851% | -12.56% | -5.915% | 2.543%
| Z-Score (Absolute) |  | 0.62 | 2.34 | 0.26 | 4.1 | 0.26 | 1.73 | 0.73 | 0.55
| Analytical results for Ni in OREAS 906b (Certified Value 3.22 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 3.1 | 3 | 2.8637 | 4 | 3.93 | 3.4 | 2.9 | 3.9 | 3.1 | 3.1 | 3 | 1.4 | 3.3 | 3.4894 | 3 | 3.2 | 2 | 2.9 | 2.8 | 4.8 | 4.1 | 4 | 3.2 | 3.1 |  | 1
| 1 | 2 | 3.3 | 3 | 2.8043 | 4.5 | 4.44 | 3 | 2.8 | 2.9 | 3.3 | 2.7 | 3 | 0.8 | 3.6 | 3.3506 | 3 | 3 | 2 | 2.6 | 3 | 4.4 | 3.8 | 3 | 3 | 3.2 |  | 11
| 1 | 3 | 3 | 3 | 2.7719 | 4.6 | 4.13 | 3.2 | 2.8 | 3.2 | 4.4 | 3.3 | 3 | 1.2 | 3.4 | 3.4589 | 3 | 3.1 | 2 | 2.6 | 2.9 | 5 | 3.9 | 3 | 3 | 3 |  | 16
| 1 | 4 | 3.2 | 3 | 2.7772 | 4.2 | 3.9 | 3.7 | 2.8 | 2.8 | 3.4 | 3 | 3 | 1.2 | 3.3 | 3.3691 | 3 | 3.2 | 2 | 2.6 | 2.7 | 4.79 | 3.8 | 2 | 3.4 | 3 |  | 3.2247
| 1 | 5 | 3.1 | 3 | 2.8884 | 4.7 | 4.03 | 3.2 | 3 | 2.8 | 4 | 3 | 3 | 1.2 | 3.2 | 3.5369 | 3 | 5.5 | 2 | 2.9 | 3 | 4.91 | 4.5 | 3 | 3.3 | 3 |  | 51
| 1 | 6 | 3.1 | 3 | 2.8716 | 5.1 | 4.19 | 3.3 | 2.8 | 2.8 | 3.6 | 3.6 | 3 | 1.2 | 3.4 | 3.3508 | 3 | 3.1 | 2 | 2.6 | 2.8 | 4.9 | 4.1 | 3 | 3 | 3.1
| Mean |  | 3.1333 | 3 | 2.8295 | 4.5167 | 4.1033 | 3.3 | 2.85 | 3.0667 | 3.6333 | 3.1167 | 3 | 1.1667 | 3.3667 | 3.426 | 3 | 3.5167 | 2 | 2.7 | 2.8667 | 4.8 | 4.0333 | 3 | 3.15 | 3.0667
| Median |  | 3.1 | 3 | 2.834 | 4.55 | 4.08 | 3.25 | 2.8 | 2.85 | 3.5 | 3.05 | 3 | 1.2 | 3.35 | 3.414 | 3 | 3.15 | 2 | 2.6 | 2.85 | 4.85 | 4 | 3 | 3.1 | 3.05
| Std Dev. |  | 0.10328 | 0 | 0.051178 | 0.38687 | 0.19916 | 0.23664 | 0.083666 | 0.43665 | 0.48442 | 0.30605 | 0 | 0.19664 | 0.13663 | 0.079949 | 0 | 0.97451 | 0 | 0.15492 | 0.12111 | 0.21081 | 0.26583 | 0.63246 | 0.17607 | 0.08165
| Rel.Std.Dev. |  | 3.296% | 0% | 1.809% | 8.565% | 4.854% | 7.171% | 2.936% | 14.24% | 13.33% | 9.82% | 0% | 16.85% | 4.058% | 2.334% | 0% | 27.71% | 0% | 5.738% | 4.225% | 4.392% | 6.591% | 21.08% | 5.589% | 2.662%
| PDM3 |  | -2.834% | -6.969% | -12.25% | 40.06% | 27.25% | 2.335% | -11.62% | -4.901% | 12.67% | -3.351% | -6.969% | -63.82% | 4.402% | 6.24% | -6.969% | 9.053% | -37.98% | -16.27% | -11.1% | 48.85% | 25.08% | -6.969% | -2.317% | -4.901%
| Z-Score (Absolute) |  | 0.04 | NA | 0.72 | 3.07 | 2.14 | 0.34 | 0.67 | 0.19 | 1.09 | 0.07 | NA | 4.46 | 0.49 | 0.62 | NA | 0.82 | NA | 1.01 | 0.64 | 3.71 | 1.99 | NA | 0 | 0.19
Results from laboratories 2.02, 2.12, 2.16, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 906b (Certified Value 0.028 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.0284 | 0.028 | 0.02697 | 0.029 | 0.05 | 0.027 | 0.028 | 0.028 | 0.029 | 0.028 | 0.029 | 0.03 | 0.0276 | 0.02878 | 0.031 | 0.0275 | 0.0272 | 0.0296 | 0.0284 | 0.0282 | 0.03 | 0.028 | 0.0274 | 0.028 |  | 1
| 1 | 2 | 0.0282 | 0.028 | 0.026034 | 0.029 | 0.089 | 0.027 | 0.027 | 0.029 | 0.03 | 0.028 | 0.029 | 0.029 | 0.0268 | 0.028741 | 0.032 | 0.0277 | 0.0282 | 0.0283 | 0.029 | 0.0286 | 0.03 | 0.024 | 0.028 | 0.028 |  | #N/A
| 1 | 3 | 0.0281 | 0.028 | 0.027707 | 0.029 | 0.084 | 0.027 | 0.027 | 0.029 | 0.029 | 0.028 | 0.029 | 0.033 | 0.027 | 0.028357 | 0.03 | 0.0277 | 0.028 | 0.0271 | 0.0279 | 0.0285 | 0.03 | 0.026 | 0.0283 | 0.026 |  | 16
| 1 | 4 | 0.0281 | 0.028 | 0.025968 | 0.029 | 0.067 | 0.027 | 0.027 | 0.029 | 0.03 | 0.027 | 0.028 | 0.0275 | 0.0263 | 0.02913 | 0.031 | 0.0275 | 0.0282 | 0.0277 | 0.0257 | 0.0273 | 0.03 | 0.025 | 0.0278 | 0.028 |  | 0.028074
| 1 | 5 | 0.0282 | 0.028 | 0.026515 | 0.029 | 0.078 | 0.028 | 0.028 | 0.029 | 0.029 | 0.028 | 0.029 | 0.028 | 0.0267 | 0.028584 | 0.031 | 0.0274 | 0.0283 | 0.0283 | 0.0286 | 0.028 | 0.03 | 0.029 | 0.028 | 0.027 |  | 52
| 1 | 6 | 0.0282 | 0.028 | 0.027496 | 0.029 | 0.085 | 0.027 | 0.028 | 0.029 | 0.029 | 0.028 | 0.029 | 0.03 | 0.0273 | 0.028814 | 0.031 | 0.0273 | 0.0281 | 0.0244 | 0.0246 | 0.0274 | 0.03 | 0.028 | 0.0282 | 0.026
| Mean |  | 0.0282 | 0.028 | 0.026782 | 0.029 | 0.0755 | 0.027167 | 0.0275 | 0.028833 | 0.029333 | 0.027833 | 0.028833 | 0.029583 | 0.02695 | 0.028734 | 0.031 | 0.027517 | 0.028 | 0.027567 | 0.027367 | 0.028 | 0.03 | 0.026667 | 0.02795 | 0.027167
| Median |  | 0.0282 | 0.028 | 0.026743 | 0.029 | 0.081 | 0.027 | 0.0275 | 0.029 | 0.029 | 0.028 | 0.029 | 0.0295 | 0.0269 | 0.02876 | 0.031 | 0.0275 | 0.02815 | 0.028 | 0.02815 | 0.0281 | 0.03 | 0.027 | 0.028 | 0.0275
| Std Dev. |  | 0.00010954 | 0 | 0.00073402 | 0 | 0.014653 | 0.00040825 | 0.00054772 | 0.00040825 | 0.0005164 | 0.00040825 | 0.00040825 | 0.00196 | 0.00045935 | 0.00025691 | 0.00063246 | 0.00016021 | 0.00040497 | 0.0017592 | 0.0017874 | 0.00054772 | 0 | 0.0019664 | 0.00032094 | 0.00098319
| Rel.Std.Dev. |  | 0.3885% | 0% | 2.741% | 0% | 19.41% | 1.503% | 1.992% | 1.416% | 1.76% | 1.467% | 1.416% | 6.625% | 1.704% | 0.8941% | 2.04% | 0.5822% | 1.446% | 6.381% | 6.531% | 1.956% | 0% | 7.374% | 1.148% | 3.619%
| PDM3 |  | 0.4491% | -0.2633% | -4.602% | 3.299% | 168.9% | -3.232% | -2.044% | 2.705% | 4.486% | -0.857% | 2.705% | 5.377% | -4.003% | 2.352% | 10.42% | -1.985% | -0.2633% | -1.807% | -2.519% | -0.2633% | 6.861% | -5.013% | -0.4414% | -3.232%
| Z-Score (Absolute) |  | 0.16 | 0 | 0.99 | 0.81 | 38.44 | 0.67 | 0.4 | 0.67 | 1.08 | 0.13 | 0.67 | 1.28 | 0.85 | 0.59 | 2.43 | 0.39 | 0 | 0.35 | 0.51 | 0 | 1.62 | 1.08 | 0.04 | 0.67
| Analytical results for Pb in OREAS 906b (Certified Value 31.7 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 32 | 31 | 32.0237 | 32.66 | 32.9 | 31.8 | 29.5 | 32.3 | 32.7 | 33.3 | 31.3 | 108 | 31.3 | 33.2422 | 33 | 31 | 31.45 | 26 | 33 | 31.4 | 36.7 | 30 | 27 | 30 | 33.88 | 1
| 1 | 2 | 32.3 | 32 | 33.1659 | 34.81 | 31.3 | 31.5 | 29.7 | 31.9 | 32.5 | 31.9 | 32.1 | 128 | 31.1 | 33.2328 | 32 | 31 | 28.32 | 27 | 32.8 | 32.5 | 36.2 | 28 | 26 | 30 | 34.24 | 29
| 1 | 3 | 31.9 | 31 | 31.4392 | 32.11 | 34.7 | 31.9 | 30.4 | 32.4 | 32.9 | 30.8 | 31.8 | 133 | 31.8 | 30.9251 | 33 | 31.2 | 27.74 | 27 | 32.5 | 32.9 | 35.6 | 30 | 23 | 31 | 32.39 | 16
| 1 | 4 | 31.5 | 32 | 31.7226 | 33.89 | 31.5 | 29.9 | 29.3 | 31.4 | 32.6 | 30.5 | 31.2 | 138 | 32 | 31.871 | 32 | 31.1 | 30.22 | 26 | 32.6 | 33.4 | 35.9 | 30 | 24 | 31 | 33.28 | 31.6988
| 1 | 5 | 31.9 | 31 | 32.8791 | 34.6 | 32 | 32.2 | 28.8 | 32.5 | 33 | 31.9 | 31.4 | 128 | 31.3 | 31.6698 | 31 | 30.6 | 30.49 | 28 | 33.5 | 31.2 | 35.2 | 30 | 28 | 31 | 33.84 | 53
| 1 | 6 | 31.6 | 32 | 32.2711 | 35.29 | 31.3 | 32 | 28.4 | 30.4 | 31.3 | 32.5 | 31.7 | 124 | 31.2 | 32.6254 | 33 | 29.8 | 26.9 | 28 | 31.1 | 33.2 | 35.1 | 29 | 26 | 29 | 32.55
| Mean |  | 31.8667 | 31.5 | 32.2503 | 33.8933 | 32.2833 | 31.55 | 29.35 | 31.8167 | 32.5 | 31.8167 | 31.5833 | 126.5 | 31.45 | 32.261 | 32.3333 | 30.7833 | 29.1867 | 27 | 32.5833 | 32.4333 | 35.7833 | 29.5 | 25.6667 | 30.3333 | 33.3633
| Median |  | 31.9 | 31.5 | 32.1474 | 34.245 | 31.75 | 31.85 | 29.4 | 32.1 | 32.65 | 31.9 | 31.55 | 128 | 31.3 | 32.2482 | 32.5 | 31 | 29.27 | 27 | 32.7 | 32.7 | 35.75 | 30 | 26 | 30.5 | 33.56
| Std Dev. |  | 0.28752 | 0.54772 | 0.66661 | 1.2643 | 1.3303 | 0.84083 | 0.70071 | 0.80353 | 0.61644 | 1.0439 | 0.34303 | 10.2713 | 0.36194 | 0.93019 | 0.8165 | 0.52313 | 1.7867 | 0.89443 | 0.8085 | 0.93095 | 0.61128 | 0.83666 | 1.8619 | 0.8165 | 0.75875
| Rel.Std.Dev. |  | 0.9023% | 1.739% | 2.067% | 3.73% | 4.121% | 2.665% | 2.387% | 2.526% | 1.897% | 3.281% | 1.086% | 8.12% | 1.151% | 2.883% | 2.525% | 1.699% | 6.122% | 3.313% | 2.481% | 2.87% | 1.708% | 2.836% | 7.254% | 2.692% | 2.274%
| PDM3 |  | 0.5295% | -0.6272% | 1.74% | 6.923% | 1.844% | -0.4695% | -7.41% | 0.3718% | 2.527% | 0.3718% | -0.3643% | 299.1% | -0.785% | 1.774% | 2.002% | -2.888% | -7.925% | -14.82% | 2.79% | 2.317% | 12.89% | -6.937% | -19.03% | -4.308% | 5.251%
| Z-Score (Absolute) |  | 0.05 | 0.31 | 0.43 | 2.05 | 0.46 | 0.26 | 2.43 | 0 | 0.67 | 0 | 0.23 | 93.43 | 0.36 | 0.44 | 0.51 | 1.02 | 2.6 | 4.75 | 0.76 | 0.61 | 3.91 | 2.29 | 6.07 | 1.46 | 1.53
| Analytical results for Pr in OREAS 906b (Certified Value 10.2 ppm) | [26 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 1
2
| 1 | 1 | 11.04 | 11.4 | 10.3462 | 7.3647 | 10.2 | 8.2 | 10.3 | [18 empty] | 1
| 1 | 2 | 10.81 | 12.5 | 10.6296 | 7.3555 | 9.85 | 10.5 | 10.7 | [18 empty] | 20
| 1 | 3 | 10.53 | 11.4 | 10.6176 | 7.3647 | 9.97 | 10.3 | 10.2 | [18 empty] | 16
| 1 | 4 | 10.55 | 12.1 | 10.6135 | 7.3474 | 9.92 | 9.4 | 10.4 | [18 empty] | 10.1972
| 1 | 5 | 10.58 | 12 | 10.4029 | 7.3217 | 10 | 7.9 | 10.3 | [18 empty] | 54
| 1 | 6 | 10.91 | 12.3 | 10.6647 | 7.3878 | 9.98 | 8.6 | 10
| Mean |  | 10.7367 | 11.95 | 10.5458 | 7.357 | 9.9867 | 9.15 | 10.3167
| Median |  | 10.695 | 12.05 | 10.6156 | 7.3601 | 9.975 | 9 | 10.3
| Std Dev. |  | 0.21426 | 0.45935 | 0.13502 | 0.021927 | 0.11759 | 1.0932 | 0.23166
| Rel.Std.Dev. |  | 1.996% | 3.844% | 1.28% | 0.298% | 1.177% | 11.95% | 2.245%
| PDM3 |  | 5.291% | 17.19% | 3.419% | -27.85% | -2.064% | -10.27% | 1.172%
| Z-Score (Absolute) |  | 0.67 | 2.62 | 0.37 | 4.75 | 0.53 | 1.87 | 0
| Analytical results for Pt in OREAS 906b (Indicative Value 11.3 ppb) | [26 empty] | Indicative
Pt | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.20 | [24 empty] | ppb
|  |  | 4A*OES/MS | [24 empty] | 1
1
| 1 | 1 | 12 | [24 empty] | 1
| 1 | 2 | 10 | [24 empty] | 1
| 1 | 3 | 12 | [24 empty] | 16
| 1 | 4 | 13 | [24 empty] | 11.3333
| 1 | 5 | 11 | [24 empty] | 13
| 1 | 6 | 10
| Mean |  | 11.3333
| Median |  | 11.5
| Std Dev. |  | 1.2111
| Rel.Std.Dev. |  | 10.69%
| PDM3 |  | 2.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 906b (Certified Value 138 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.20 | 2.23 | 2.24 | 2.25 | 2.27 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
0
| 1 | 1 | 143.47 | 133 | 138.5713 | 186.4 | 130.01 | 150 | 133.5 | 144 | 140.5 | 130.5 | 138 | 106 | 131.2 | 152.213 | 105.3378 | 135 | 138.43 | 109 | 150 | 144.5 | [5 empty] | 1
| 1 | 2 | 140.67 | 133 | 139.3379 | 183.4 | 125.67 | 146.5 | 131.5 | 143 | 142 | 124.5 | 142 | 105 | 130.7 | 150.5571 | 105.375 | 134 | 136.94 | 103 | 106 | 136.7 | [5 empty] | 45
| 1 | 3 | 141.84 | 131 | 138.489 | 169.2 | 128.83 | 150 | 132 | 146.5 | 142.5 | 137.5 | 138 | 104 | 134.9 | 151.7698 | 105.3278 | 135 | 137.43 | 106 | 128 | 139.4 | [5 empty] | 16
| 1 | 4 | 141.78 | 136 | 138.0627 | 182.8 | 125.8 | 145 | 133.5 | 147 | 145.5 | 124 | 137 | 107 | 131.2 | 148.0784 | 105.3964 | 135 | 138.83 | 108 | 163 | 136.8 | [5 empty] | 138.48
| 1 | 5 | 139.88 | 133 | 139.4649 | 171.4 | 128.65 | 149.5 | 135.5 | 145 | 143 | 130.5 | 137 | 102 | 131.8 | 145.962 | 105.3864 | 140 | 139.45 | 103 | 150 | 138.1 | [5 empty] | 55
| 1 | 6 | 142.92 | 132 | 139.0677 | 147.6 | 132.17 | 150 | 129 | 144.5 | 139.5 | 130.5 | 135 | 104 | 134.6 | 150.6231 | 105.3168 | 135 | 139.31 | 119 | 146 | 134.4
| Mean |  | 141.76 | 133 | 138.8323 | 173.4667 | 128.5217 | 148.5 | 132.5 | 145 | 142.1667 | 129.5833 | 137.8333 | 104.6667 | 132.4 | 149.8672 | 105.3567 | 135.6667 | 138.3983 | 108 | 140.5 | 138.3167
| Median |  | 141.81 | 133 | 138.8195 | 177.1 | 128.74 | 149.75 | 132.75 | 144.75 | 142.25 | 130.5 | 137.5 | 104.5 | 131.5 | 150.5901 | 105.3564 | 135 | 138.63 | 107 | 148 | 137.45
| Std Dev. |  | 1.3416 | 1.6733 | 0.54581 | 14.4542 | 2.4971 | 2.1909 | 2.2136 | 1.5166 | 2.0897 | 4.9439 | 2.3166 | 1.7512 | 1.8558 | 2.3925 | 0.033399 | 2.1602 | 1.0187 | 5.933 | 20.315 | 3.4557
| Rel.Std.Dev. |  | 0.9464% | 1.258% | 0.3931% | 8.333% | 1.943% | 1.475% | 1.671% | 1.046% | 1.47% | 3.815% | 1.681% | 1.673% | 1.402% | 1.596% | 0.0317% | 1.592% | 0.7361% | 5.493% | 14.46% | 2.498%
| PDM3 |  | 2.369% | -3.957% | 0.2544% | 25.26% | -7.191% | 7.236% | -4.318% | 4.708% | 2.662% | -6.424% | -0.467% | -24.42% | -4.391% | 8.223% | -23.92% | -2.032% | -0.05895% | -22.01% | 1.459% | -0.1179%
| Z-Score (Absolute) |  | 0.44 | 0.61 | 0.09 | 4.24 | 1.15 | 1.25 | 0.67 | 0.83 | 0.49 | 1.02 | 0.03 | 4 | 0.68 | 1.41 | 3.92 | 0.29 | 0.04 | 3.61 | 0.29 | 0.03
| Analytical results for Re in OREAS 906b (Certified Value 0.002 ppm) | [26 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.15 | 2.20 | 2.25 | 2.26 | 2.27 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 |  | < 0.05 | < 0.05 | 0.019 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | 0.063054 | 0.0023 | 0.001 | NA | < 0.002 | [10 empty] | 1
| 1 | 2 |  | < 0.05 | < 0.05 | 0.02 | 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.003 | 0.002 | 0.073605 | 0.0025 | 0.001 | NA | < 0.002 | [10 empty] | 12
| 1 | 3 | 0.003 | < 0.05 | < 0.05 | 0.015 | 0.002 | 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | 0.10208 | 0.0023 | 0.001 | NA | < 0.002 | [10 empty] | 16
| 1 | 4 |  | < 0.05 | < 0.05 | 0.016 | 0.002 | 0.002 | < 0.002 | 0.002 | 0.003 | 0.003 | 0.11462 | 0.0024 | 0.002 | NA | < 0.002 | [10 empty] | 0.0022146
| 1 | 5 |  | < 0.05 | < 0.05 | 0.026 | 0.003 | < 0.002 | < 0.002 | < 0.002 | 0.003 | 0.003 | 0.065462 | 0.002 | 0.001 | NA | < 0.002 | [10 empty] | 56
| 1 | 6 |  | < 0.05 | < 0.05 | 0.008 | 0.002 | 0.002 | < 0.002 | < 0.002 | 0.002 | 0.003 | 0.07311 | 0.002 | 0.001 | NA | < 0.002
| Mean |  | 0.003 |  |  | 0.017333 | 0.0022 | 0.002 |  | 0.002 | 0.0026 | 0.0025 | 0.081989 | 0.00225 | 0.0011667
| Median |  | 0.003 |  |  | 0.0175 | 0.002 | 0.002 |  | 0.002 | 0.003 | 0.0025 | 0.073357 | 0.0023 | 0.001
| Std Dev. |  |  |  |  | 0.0059889 | 0.00044721 | 0 |  | 0 | 0.00054772 | 0.00054772 | 0.02121 | 0.00020736 | 0.00040825
| Rel.Std.Dev. |  |  |  |  | 34.55% | 20.33% | 0% |  | 0% | 21.07% | 21.91% | 25.87% | 9.216% | 34.99%
| PDM3 |  | 35.47% |  |  | 682.7% | -0.6585% | -9.69% |  | -9.69% | 17.4% | 12.89% | 3602% | 1.599% | -47.32%
| Z-Score (Absolute) |  | 0.46 | 14.7 | 14.7 | 9.74 | 0.19 | 0.4 | 0.84 | 0.62 | 0.03 | 0.13 | 51.59 | 0.03 | 0.73 | NA | 0.84
| Analytical results for S in OREAS 906b (Certified Value 0.044 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.041 | 0.04 | 0.039743 | 0.04 | 0.03 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.044 | 0.05 | 0.04 | 0.052316 | 0.047 | 0.04 | 0.045442 | 0.04 | 0.0327 | 0.06 | 0.0444 | 0.05 | 0.04 | 0.0413 | 0.04 | 1
| 1 | 2 | 0.0404 | 0.041 | 0.040023 | < 0.04 | 0.03 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.044 | 0.05 | 0.04 | 0.051202 | 0.049 | 0.04 | 0.046256 | 0.04 | 0.0322 | 0.06 | 0.0456 | 0.05 | 0.04 | 0.0418 | 0.04 | 3
| 1 | 3 | 0.0408 | 0.042 | 0.039785 | 0.04 | 0.02 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.044 | 0.05 | 0.04 | 0.054258 | 0.046 | 0.04 | 0.045656 | 0.04 | 0.039 | 0.05 | 0.0444 | 0.05 | 0.04 | 0.0424 | < 0.04 | 16
| 1 | 4 | 0.04 | 0.041 | 0.039237 | 0.04 | 0.02 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.044 | 0.047 | 0.03 | 0.052168 | 0.048 | 0.04 | 0.048712 | 0.04 | 0.0283 | 0.05 | 0.0465 | 0.05 | 0.04 | 0.0419 | < 0.04 | 0.044263
| 1 | 5 | 0.0405 | 0.041 | 0.039248 | 0.04 | 0.03 | 0.05 | 0.05 | 0.04 | 0.05 | 0.04 | 0.044 | 0.05 | 0.04 | 0.052929 | 0.045 | 0.04 | 0.044154 | 0.04 | 0.0377 | 0.06 | 0.0466 | 0.05 | 0.04 | 0.041 | < 0.04 | 57
| 1 | 6 | 0.0405 | 0.04 | 0.040044 | 0.04 | 0.03 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.043 | 0.05 | 0.04 | 0.05371 | 0.047 | 0.04 | 0.045739 | 0.04 | 0.0314 | 0.05 | 0.0452 | 0.05 | 0.04 | 0.0409 | < 0.04
| Mean |  | 0.040533 | 0.040833 | 0.03968 | 0.04 | 0.026667 | 0.048333 | 0.05 | 0.048333 | 0.05 | 0.041667 | 0.043833 | 0.0495 | 0.038333 | 0.052764 | 0.047 | 0.04 | 0.045993 | 0.04 | 0.03355 | 0.055 | 0.04545 | 0.05 | 0.04 | 0.04155 | 0.04
| Median |  | 0.0405 | 0.041 | 0.039764 | 0.04 | 0.03 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.044 | 0.05 | 0.04 | 0.052623 | 0.047 | 0.04 | 0.045698 | 0.04 | 0.03245 | 0.055 | 0.0454 | 0.05 | 0.04 | 0.04155 | 0.04
| Std Dev. |  | 0.00034448 | 0.00075277 | 0.00035998 | 0 | 0.005164 | 0.0040825 | 7.6012e-18 | 0.0040825 | 7.6012e-18 | 0.0040825 | 0.00040825 | 0.0012247 | 0.0040825 | 0.0011091 | 0.0014142 | 0 | 0.001505 | 0 | 0.0040412 | 0.0054772 | 0.00097108 | 7.6012e-18 | 0 | 0.00058224 | 0
| Rel.Std.Dev. |  | 0.8499% | 1.844% | 0.9072% | 0% | 19.36% | 8.447% | 1.52e-14% | 8.447% | 1.52e-14% | 9.798% | 0.9314% | 2.474% | 10.65% | 2.102% | 3.009% | 0% | 3.272% | 0% | 12.05% | 9.959% | 2.137% | 1.52e-14% | 0% | 1.401% | 0%
| PDM3 |  | -8.426% | -7.748% | -10.35% | -9.631% | -39.75% | 9.196% | 12.96% | 9.196% | 12.96% | -5.865% | -0.9705% | 11.83% | -13.4% | 19.21% | 6.184% | -9.631% | 3.909% | -9.631% | -24.2% | 24.26% | 2.682% | 12.96% | -9.631% | -6.129% | -9.631%
| Z-Score (Absolute) |  | 0.15 | 0.11 | 0.27 | 0.67 | 2.02 | 0.9 | 1.12 | 0.9 | 1.12 | 0 | 0.29 | 1.06 | 0.45 | 1.5 | 0.72 | 0.22 | 0.58 | 0.22 | 1.09 | 1.8 | 0.51 | 1.12 | 0.22 | 0.02 | 2.02
| Analytical results for Sb in OREAS 906b (Certified Value 2.16 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.21 | 2.1 | 2.3814 | 2.07 | 1.22 | 2.17 | 2.12 | 2.36 | 2.47 | 2.1 | 2.1 | 4.88 | 2.2 | 2.1534 | < 5 | 2.14 | < 5 | 2.1 | 10.49 | 2.5 | 2.19 | 2.4 | 2.07 | 1.93 |  | 1
| 1 | 2 | 2.17 | 2.1 | 2.4657 | 2.22 | 1.23 | 2.17 | 2.05 | 2.1 | 2.51 | 2.01 | 2.1 | 2.93 | 2.06 | 2.1782 | < 5 | 2.08 | < 5 | 2.1 | 10.98 | 2.4 | 2.29 | 2 | 2 | 1.88 |  | 47
| 1 | 3 | 2.08 | 2.1 | 2.4327 | 2.04 | 1.22 | 2.19 | 2.07 | 2.17 | 2.55 | 1.98 | 2.1 | 5.61 | 2.1 | 2.1394 | < 5 | 2.11 | < 5 | 2.1 | 11.14 | 2.35 | 2.24 | 2.1 | 2 | 1.82 |  | 16
| 1 | 4 | 2.15 | 2.1 | 2.4081 | 2.15 | 1.2 | 2.19 | 2.08 | 2.24 | 2.46 | 2.08 | 2.05 | 3.11 | 2.07 | 2.1396 | < 5 | 2.16 | < 5 | 2.1 | 10.9 | 2.4 | 2.33 | 2.4 | 2.08 | 1.85 |  | 2.1562
| 1 | 5 | 2.17 | 2 | 2.364 | 2.25 | 1.19 | 2.22 | 2.09 | 2.37 | 2.56 | 2.15 | 2 | 2.51 | 2.12 | 2.1376 | < 5 | 2.21 | < 5 | 2.1 | 11.96 | 2.3 | 2.25 | 2.4 | 2.14 | 1.87 |  | 58
| 1 | 6 | 2.09 | 2 | 2.4905 | 2.22 | 1.21 | 2.2 | 2.02 | 2.23 | 2.37 | 2.16 | 2.04 | 8.07 | 2.07 | 2.142 | < 5 | 2.11 | < 5 | 2.1 | 11.12 | 2.3 | 2.34 | 2.4 | 1.97 | 1.85
| Mean |  | 2.145 | 2.0667 | 2.4237 | 2.1583 | 1.2117 | 2.19 | 2.0717 | 2.245 | 2.4867 | 2.08 | 2.065 | 4.5183 | 2.1033 | 2.1484 |  | 2.135 |  | 2.1 | 11.0983 | 2.375 | 2.2733 | 2.2833 | 2.0433 | 1.8667
| Median |  | 2.16 | 2.1 | 2.4204 | 2.185 | 1.215 | 2.19 | 2.075 | 2.235 | 2.49 | 2.09 | 2.075 | 3.995 | 2.085 | 2.1408 |  | 2.125 |  | 2.1 | 11.05 | 2.375 | 2.27 | 2.4 | 2.035 | 1.86
| Std Dev. |  | 0.050498 | 0.05164 | 0.048803 | 0.087044 | 0.01472 | 0.018974 | 0.034303 | 0.10559 | 0.070048 | 0.072938 | 0.041833 | 2.1202 | 0.052409 | 0.015664 |  | 0.045935 |  | 0 | 0.48334 | 0.075829 | 0.057504 | 0.18348 | 0.064083 | 0.037238
| Rel.Std.Dev. |  | 2.354% | 2.499% | 2.014% | 4.033% | 1.215% | 0.8664% | 1.656% | 4.703% | 2.817% | 3.507% | 2.026% | 46.92% | 2.492% | 0.7291% |  | 2.152% |  | 0% | 4.355% | 3.193% | 2.529% | 8.036% | 3.136% | 1.995%
| PDM3 |  | -0.5182% | -4.151% | 12.41% | 0.1002% | -43.8% | 1.569% | -3.919% | 4.12% | 15.33% | -3.533% | -4.228% | 109.6% | -2.451% | -0.3616% |  | -0.982% |  | -2.605% | 414.7% | 10.15% | 5.434% | 5.897% | -5.233% | -13.43%
| Z-Score (Absolute) |  | 0.06 | 0.58 | 1.81 | 0.03 | 6.3 | 0.25 | 0.55 | 0.61 | 2.23 | 0.49 | 0.59 | 15.82 | 0.33 | 0.03 | 2.32 | 0.12 | 2.32 | 0.36 | 59.82 | 1.48 | 0.8 | 0.87 | 0.74 | 1.92
| Analytical results for Sc in OREAS 906b (Certified Value 4.71 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 4.8 | 4 | 4.9176 | 6.4 | 1.13 | 4.7 | 4.2 | 4.5 | 5 | 4.2 | 4.4 | 5.4 | 4.3 | 4.8047 | 4.7 | 5.7 | 3.5603 | < 5 | 6 | 6.1 | 4.98 | 4.9 | 4 | 4.6 |  | 1
| 1 | 2 | 4.3 | 4 | 5.2131 | 7 | 1.15 | 4.5 | 4 | 4.7 | 5.2 | 4 | 4.4 | 3.5 | 4 | 4.9182 | 4.9 | 5.5 | 3.5946 | < 5 | 5 | 5.6 | 4.68 | 4.5 | 5 | 4.5 |  | 14
| 1 | 3 | 4.6 | 4 | 5.1096 | 6.5 | 1.13 | 4.6 | 4.1 | 4.8 | 5.2 | 4.5 | 4.4 | 5.2 | 4.3 | 4.7544 | 4.5 | 5.2 | 3.5613 | < 5 | 5 | 6 | 5.03 | 4.4 | 5 | 4.3 |  | 16
| 1 | 4 | 4.7 | 4 | 4.9377 | 6.9 | 1.14 | 4.5 | 4 | 4.7 | 5.1 | 4 | 4.3 | 5.03 | 4.3 | 4.8979 | 4.6 | 5.5 | 3.5831 | < 5 | 5 | 5.6 | 5 | 4.7 | 5 | 4.5 |  | 4.7135
| 1 | 5 | 4.7 | 4 | 4.9554 | 6.7 | 1.13 | 4.6 | 4.1 | 4.5 | 5.1 | 4.2 | 4.3 | 5.4 | 4.4 | 4.803 | 4.7 | 5.5 | 3.5629 | < 5 | 5 | 6.1 | 4.98 | 4.9 | 4 | 4.5 |  | 59
| 1 | 6 | 4.9 | 4 | 5.2024 | 6.9 | 1.13 | 4.7 | 3.9 | 4.7 | 5 | 4.2 | 4.3 | 5.3 | 4.5 | 4.94 | 5 | 5.6 | 3.5549 | < 5 | 5 | 5.4 | 5 | 4.8 | 5 | 4.4
| Mean |  | 4.6667 | 4 | 5.0559 | 6.7333 | 1.135 | 4.6 | 4.05 | 4.65 | 5.1 | 4.1833 | 4.35 | 4.9717 | 4.3 | 4.853 | 4.7333 | 5.5 | 3.5695 |  | 5.1667 | 5.8 | 4.945 | 4.7 | 4.6667 | 4.4667
| Median |  | 4.7 | 4 | 5.0325 | 6.8 | 1.13 | 4.6 | 4.05 | 4.7 | 5.1 | 4.2 | 4.35 | 5.25 | 4.3 | 4.8513 | 4.7 | 5.5 | 3.5621 |  | 5 | 5.8 | 4.99 | 4.75 | 5 | 4.5
| Std Dev. |  | 0.20656 | 0 | 0.13584 | 0.24221 | 0.0083666 | 0.089443 | 0.10488 | 0.12247 | 0.089443 | 0.18348 | 0.054772 | 0.73431 | 0.16733 | 0.075356 | 0.18619 | 0.16733 | 0.015616 |  | 0.40825 | 0.30332 | 0.13111 | 0.20976 | 0.5164 | 0.10328
| Rel.Std.Dev. |  | 4.426% | 0% | 2.687% | 3.597% | 0.7371% | 1.944% | 2.59% | 2.634% | 1.754% | 4.386% | 1.259% | 14.77% | 3.891% | 1.553% | 3.934% | 3.042% | 0.4375% |  | 7.902% | 5.23% | 2.651% | 4.463% | 11.07% | 2.312%
| PDM3 |  | -0.993% | -15.14% | 7.266% | 42.85% | -75.92% | -2.407% | -14.08% | -1.347% | 8.201% | -11.25% | -7.711% | 5.478% | -8.772% | 2.961% | 0.4214% | 16.69% | -24.27% |  | 9.615% | 23.05% | 4.912% | -0.2858% | -0.993% | -5.236%
| Z-Score (Absolute) |  | 0 | NA | 0.72 | 3.8 | 6.49 | 0.12 | 1.13 | 0.03 | 0.8 | 0.89 | 0.58 | 0.56 | 0.67 | 0.34 | 0.12 | 1.53 | 2.02 | 3.98 | NA | 2.08 | 0.51 | 0.06 | NA | 0.37
Results from laboratories 2.02, 2.20 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 906b (Certified Value 5.11 ppm) | [26 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.1 | 5 | 4.9212 | 5 | 4.24 | 5 | 6 | 5 | 7 | 5 | 4 | 5 | 6 | 5.4843 | 4.8 | < 10 | 4.1 | 3 | 4.74 | < 2 | 5 | 4.9 | 5.5 |  |  | 1
| 1 | 2 | 5 | 5 | 5.1037 | 5.7 | 4.21 | 5 | 5 | 5 | 6 | 4 | 5 | 5 | 7 | 5.4005 | 5.5 | < 10 | 3.73 | 4 | 4.92 | 5 | 4.5 | 4.85 | 5.3 |  |  | 49
| 1 | 3 | 5.3 | 5 | 5.0697 | 4.7 | 4.22 | 6 | 6 | 6 | 6 | 5 | 5 | 6 | 6 | 5.7542 | 5.5 | < 10 | 3.97 | 3 | 4.84 | 6 | 5.1 | 4.85 | 5.2 |  |  | 16
| 1 | 4 | 4.8 | 6 | 4.9381 | 5.9 | 4.22 | 4 | 5 | 5 | 7 | 4 | 4 | 5.2 | 6 | 5.6974 | 5.3 | < 10 | 4.42 | 4 | 4.83 | 3 | 5.3 | 4.96 | 5.1 |  |  | 5.1088
| 1 | 5 | 4.6 | 5 | 5.067 | 5.7 | 4.22 | 5 | 6 | 5 | 6 | 5 | 6 | 5 | 6 | 5.7803 | 5.2 | < 10 | 3.53 | 4 | 4.92 | 7 | 5.2 | 4.96 | 5 |  |  | 60
| 1 | 6 | 4.9 | 5 | 5.0825 | 5.3 | 4.23 | 6 | 5 | 5 | 6 | 5 | 5 | 5 | 5 | 5.4452 | 4.5 | < 10 | 4.09 | 4 | 4.58 | 7 | 4.9 | 4.73 | 5.3
| Mean |  | 4.95 | 5.1667 | 5.0304 | 5.3833 | 4.2233 | 5.1667 | 5.5 | 5.1667 | 6.3333 | 4.6667 | 4.8333 | 5.2 | 6 | 5.5937 | 5.1333 |  | 3.9733 | 3.6667 | 4.805 | 5.6 | 5 | 4.875 | 5.2333
| Median |  | 4.95 | 5 | 5.0684 | 5.5 | 4.22 | 5 | 5.5 | 5 | 6 | 5 | 5 | 5 | 6 | 5.5909 | 5.25 |  | 4.03 | 4 | 4.835 | 6 | 5.05 | 4.875 | 5.25
| Std Dev. |  | 0.2429 | 0.40825 | 0.079284 | 0.46655 | 0.010328 | 0.75277 | 0.54772 | 0.40825 | 0.5164 | 0.5164 | 0.75277 | 0.4 | 0.63246 | 0.16892 | 0.40332 |  | 0.31143 | 0.5164 | 0.12896 | 1.6733 | 0.28284 | 0.086429 | 0.17512
| Rel.Std.Dev. |  | 4.907% | 7.902% | 1.576% | 8.667% | 0.2445% | 14.57% | 9.959% | 7.902% | 8.154% | 11.07% | 15.57% | 7.692% | 10.54% | 3.02% | 7.857% |  | 7.838% | 14.08% | 2.684% | 29.88% | 5.657% | 1.773% | 3.346%
| PDM3 |  | -3.109% | 1.133% | -1.535% | 5.374% | -17.33% | 1.133% | 7.657% | 1.133% | 23.97% | -8.655% | -5.392% | 1.785% | 17.44% | 9.49% | 0.48% |  | -22.23% | -28.23% | -5.947% | 9.615% | -2.13% | -4.577% | 2.437%
| Z-Score (Absolute) |  | 0.28 | 0.47 | 0 | 1.21 | 2.76 | 0.47 | 1.61 | 0.47 | 4.46 | 1.24 | 0.67 | 0.58 | 3.32 | 1.93 | 0.35 | 0.1 | 3.62 | 4.67 | 0.77 | 0.67 | 0.1 | 0.53 | 0.69
| Analytical results for Sm in OREAS 906b (Certified Value 7.37 ppm) | [26 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
3
| 1 | 1 | 7.48 | 8.5 | 7.6264 | 5.3197 | 7.53 | 5.9 | 7.6 | [18 empty] | 1
| 1 | 2 | 7.26 | 9 | 7.8606 | 5.4353 | 7.21 | 8.2 | 7.6 | [18 empty] | 18
| 1 | 3 | 7.25 | 8.2 | 7.6561 | 5.2899 | 7.15 | 8.3 | 7.1 | [18 empty] | 16
| 1 | 4 | 7.23 | 8.9 | 7.6263 | 5.439 | 7.26 | 6.8 | 7.4 | [18 empty] | 7.3708
| 1 | 5 | 7.22 | 8.9 | 7.6683 | 5.3025 | 7.32 | 6.2 | 7.5 | [18 empty] | 61
| 1 | 6 | 7.53 | 9 | 7.8163 | 5.5028 | 7.61 | 6.3 | 7.5
| Mean |  | 7.3283 | 8.75 | 7.709 | 5.3815 | 7.3467 | 6.95 | 7.45
| Median |  | 7.255 | 8.9 | 7.6622 | 5.3775 | 7.29 | 6.55 | 7.5
| Std Dev. |  | 0.13848 | 0.32711 | 0.10259 | 0.088738 | 0.18359 | 1.0483 | 0.18708
| Rel.Std.Dev. |  | 1.89% | 3.738% | 1.331% | 1.649% | 2.499% | 15.08% | 2.511%
| PDM3 |  | -0.5762% | 18.71% | 4.589% | -26.99% | -0.3275% | -5.709% | 1.074%
| Z-Score (Absolute) |  | 0.03 | 2.61 | 0.67 | 3.66 | 0 | 0.74 | 0.19
| Analytical results for Sn in OREAS 906b (Certified Value 4.24 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 4.4 | 4.2 | 4.5437 | 4.7 | 2.32 | 4.6 | 4.1 | 4.4 | 4.5 | 4.1 | 4 | 3.6 | 4.2 | 4.1679 | < 10 | 4.2 | < 20 | 4 | 6.8 | 4.59 | 4.3 | 5 | 4 |  |  | 1
| 1 | 2 | 4.5 | 4.1 | 4.2771 | 5.3 | 2.32 | 4.5 | 4.1 | 4.1 | 4.6 | 3.9 | 4.2 | 1.1 | 4 | 4.2721 | < 10 | 4 | < 20 | 4.1 | 6.7 | 4.77 | 4.3 | 4 | 3.8 |  |  | 10
| 1 | 3 | 4.2 | 4.1 | 4.2422 | 5.4 | 2.31 | 4.8 | 4.2 | 4.1 | 4.6 | 3.9 | 4.1 | 2 | 4.1 | 4.4323 | < 10 | 4.1 | < 20 | 4.3 | 6.8 | 4.77 | 3.8 | 4 | 3.7 |  |  | 16
| 1 | 4 | 4.3 | 4.2 | 4.1853 | 4.5 | 2.33 | 4.4 | 4 | 4 | 4.5 | 3.9 | 4.2 | 4.6 | 4.5 | 4.4286 | < 10 | 4.2 | < 20 | 4.3 | 5.6 | 4.68 | 4.1 | 5 | 3.8 |  |  | 4.2394
| 1 | 5 | 4.5 | 4.1 | 4.4292 | 4.6 | 2.33 | 4.9 | 4.1 | 4.3 | 4.6 | 4.1 | 3.9 | 2.2 | 4.1 | 4.3035 | < 10 | 4.1 | < 20 | 4.4 | 7.2 | 4.23 | 4.3 | 5 | 3.7 |  |  | 62
| 1 | 6 | 4.3 | 4.2 | 4.4829 | 5.3 | 2.32 | 4.6 | 3.9 | 4 | 4.4 | 4.1 | 4 | 1.3 | 4.5 | 4.382 | < 10 | 4.1 | < 20 | 4.3 | 6.4 | 4.14 | 4.3 | 5 | 3.8
| Mean |  | 4.3667 | 4.15 | 4.3601 | 4.9667 | 2.3217 | 4.6333 | 4.0667 | 4.15 | 4.5333 | 4 | 4.0667 | 2.4667 | 4.2333 | 4.3311 |  | 4.1167 |  | 4.2333 | 6.5833 | 4.53 | 4.1833 | 4.6667 | 3.8
| Median |  | 4.35 | 4.15 | 4.3531 | 5 | 2.32 | 4.6 | 4.1 | 4.1 | 4.55 | 4 | 4.05 | 2.1 | 4.15 | 4.3428 |  | 4.1 |  | 4.3 | 6.75 | 4.635 | 4.3 | 5 | 3.8
| Std Dev. |  | 0.12111 | 0.054772 | 0.14485 | 0.40825 | 0.0075277 | 0.18619 | 0.10328 | 0.16432 | 0.08165 | 0.10954 | 0.12111 | 1.3677 | 0.21602 | 0.10318 |  | 0.075277 |  | 0.15055 | 0.54559 | 0.27691 | 0.20412 | 0.5164 | 0.10954
| Rel.Std.Dev. |  | 2.773% | 1.32% | 3.322% | 8.22% | 0.3242% | 4.018% | 2.54% | 3.959% | 1.801% | 2.739% | 2.978% | 55.45% | 5.103% | 2.382% |  | 1.829% |  | 3.556% | 8.287% | 6.113% | 4.879% | 11.07% | 2.883%
| PDM3 |  | 3.001% | -2.11% | 2.845% | 17.15% | -45.24% | 9.291% | -4.076% | -2.11% | 6.932% | -5.648% | -4.076% | -41.82% | -0.1442% | 2.162% |  | -2.896% |  | -0.1442% | 55.29% | 6.854% | -1.324% | 10.08% | -10.37%
| Z-Score (Absolute) |  | 0.45 | 0.28 | 0.43 | 2.47 | 6.45 | 1.35 | 0.56 | 0.28 | 1.01 | 0.79 | 0.56 | 5.96 | 0 | 0.33 | 2.58 | 0.39 | 19.44 | 0 | 7.92 | 1 | 0.17 | NA | 1.46
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 906b (Certified Value 165 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 169.65 | 170 | 173.8251 | 172 | 191.85 | 166.5 | 165 | 161.5 | 161 | 160.5 | 161 | 132 | 167.9 | 163.5912 | 168 | 159.4 | 173.8801 | 158 | 161 | 164 | 161 | 128 | 165 | 171 |  | 1
| 1 | 2 | 167.14 | 171 | 172.0903 | 173 | 190.02 | 166.5 | 157.5 | 167.5 | 161 | 157 | 160 | 126 | 166.4 | 163.0021 | 170 | 155.7 | 173.3445 | 164 | 146 | 174 | 155 | 126 | 177 | 165 |  | 25
| 1 | 3 | 167.79 | 169 | 172.0876 | 172 | 196.46 | 168.5 | 162.5 | 169.5 | 160 | 161.5 | 161 | 124 | 162.2 | 163.2447 | 165 | 157.1 | 172.5411 | 161 | 160 | 166 | 159 | 126 | 179 | 162 |  | 16
| 1 | 4 | 167.18 | 170 | 174.8701 | 171 | 183.43 | 165 | 159 | 170 | 161.5 | 156 | 158 | 126 | 167.8 | 161.7274 | 167 | 157 | 174.4466 | 162 | 150 | 162 | 158 | 128 | 176 | 165 |  | 164.7968
| 1 | 5 | 170.97 | 169 | 177.1529 | 171 | 197.72 | 170 | 162.5 | 163 | 161.5 | 160 | 160 | 128 | 166.4 | 161.6565 | 165 | 159.9 | 176.5684 | 163 | 156 | 170 | 158 | 125 | 167 | 162 |  | 63
| 1 | 6 | 168.24 | 169 | 177.3712 | 171 | 186.66 | 164 | 158 | 166.5 | 158.5 | 157.5 | 160 | 131 | 169.3 | 163.3884 | 165 | 156.1 | 172.4381 | 162 | 148 | 160 | 157 | 128 | 170 | 160
| Mean |  | 168.495 | 169.6667 | 174.5662 | 171.6667 | 191.0233 | 166.75 | 160.75 | 166.3333 | 160.5833 | 158.75 | 160 | 127.8333 | 166.6667 | 162.7684 | 166.6667 | 157.5333 | 173.8698 | 161.6667 | 153.5 | 166 | 158 | 126.8333 | 172.3333 | 164.1667
| Median |  | 168.015 | 169.5 | 174.3476 | 171.5 | 190.935 | 166.5 | 160.75 | 167 | 161 | 158.75 | 160 | 127 | 167.1 | 163.1234 | 166 | 157.05 | 173.6123 | 162 | 153 | 165 | 158 | 127 | 173 | 163.5
| Std Dev. |  | 1.5223 | 0.8165 | 2.3441 | 0.8165 | 5.5279 | 2.2079 | 3.0125 | 3.4448 | 1.1583 | 2.2079 | 1.0954 | 3.1252 | 2.4427 | 0.8559 | 2.0656 | 1.7305 | 1.5301 | 2.0656 | 6.3797 | 5.2154 | 2 | 1.3292 | 5.785 | 3.8687
| Rel.Std.Dev. |  | 0.9035% | 0.4812% | 1.343% | 0.4756% | 2.894% | 1.324% | 1.874% | 2.071% | 0.7213% | 1.391% | 0.6847% | 2.445% | 1.466% | 0.5258% | 1.239% | 1.099% | 0.88% | 1.278% | 4.156% | 3.142% | 1.266% | 1.048% | 3.357% | 2.357%
| PDM3 |  | 2.244% | 2.955% | 5.928% | 4.169% | 15.91% | 1.185% | -2.456% | 0.9324% | -2.557% | -3.669% | -2.911% | -22.43% | 1.135% | -1.231% | 1.135% | -4.408% | 5.506% | -1.899% | -6.855% | 0.7301% | -4.124% | -23.04% | 4.573% | -0.3824%
| Z-Score (Absolute) |  | 0.48 | 0.64 | 1.32 | 0.92 | 3.62 | 0.23 | 0.6 | 0.17 | 0.63 | 0.88 | 0.71 | 5.2 | 0.22 | 0.32 | 0.22 | 1.05 | 1.23 | 0.48 | 1.62 | 0.13 | 0.99 | 5.34 | 1.01 | 0.13
| Analytical results for Ta in OREAS 906b (Certified Value 1.37 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.45 | 1.3 | 1.5188 | 1.6 | 1.53 | 1.33 | 1.32 | 1.33 | 1.38 | 1.29 | 1.2 | 1.45 | 1.5 | 1.5292 | < 10 | 1.52 | 1.08 | 1.29 | 1.2 | 1.3 | 1.3 |  |  |  |  | 1
| 1 | 2 | 1.46 | 1.24 | 1.4805 | 1.9 | 1.38 | 1.31 | 1.32 | 1.35 | 1.39 | 1.23 | 1.23 | 1.57 | 1.52 | 1.4763 | < 10 | 1.5 | 1.04 | 1.3 | 1.33 | 0.4 | 1.3 |  |  |  |  | 36
| 1 | 3 | 1.38 | 1.28 | 1.5726 | 1.8 | 1.71 | 1.35 | 1.33 | 1.4 | 1.37 | 1.25 | 1.29 | 1.3 | 1.55 | 1.5843 | < 10 | 1.44 | 1.08 | 1.32 | 1.26 | 1 | 1.2 |  |  |  |  | 16
| 1 | 4 | 1.4 | 1.26 | 1.4531 | 2.3 | 1.48 | 1.32 | 1.28 | 1.38 | 1.37 | 1.24 | 1.25 | 1.31 | 1.62 | 1.5636 | < 10 | 1.46 | 0.94 | 1.27 | 1.27 | 1.3 | 1.3 |  |  |  |  | 1.3654
| 1 | 5 | 1.46 | 1.29 | 1.6157 | 2.5 | 1.52 | 1.36 | 1.32 | 1.32 | 1.38 | 1.31 | 1.25 | 1.34 | 1.59 | 1.5133 | < 10 | 1.48 | 1.07 | 1.28 | 1.21 | 1.4 | 1.3 |  |  |  |  | 64
| 1 | 6 | 1.45 | 1.33 | 1.574 | 2 | 1.31 | 1.35 | 1.26 | 1.34 | 1.36 | 1.3 | 1.27 | 1.38 | 1.49 | 1.5656 | < 10 | 1.5 | 1.07 | 1.36 | 1.19 | 1.3 | 1.3
| Mean |  | 1.4333 | 1.2833 | 1.5358 | 2.0167 | 1.4883 | 1.3367 | 1.305 | 1.3533 | 1.375 | 1.27 | 1.2483 | 1.3917 | 1.545 | 1.5387 |  | 1.4833 | 1.0467 | 1.3033 | 1.2433 | 1.1167 | 1.2833
| Median |  | 1.45 | 1.285 | 1.5457 | 1.95 | 1.5 | 1.34 | 1.32 | 1.345 | 1.375 | 1.27 | 1.25 | 1.36 | 1.535 | 1.5464 |  | 1.49 | 1.07 | 1.295 | 1.235 | 1.3 | 1.3
| Std Dev. |  | 0.034448 | 0.031411 | 0.062253 | 0.33116 | 0.13819 | 0.019664 | 0.028107 | 0.030768 | 0.010488 | 0.034059 | 0.031252 | 0.10304 | 0.051672 | 0.040143 |  | 0.029439 | 0.054283 | 0.03266 | 0.053541 | 0.37639 | 0.040825
| Rel.Std.Dev. |  | 2.403% | 2.448% | 4.054% | 16.42% | 9.285% | 1.471% | 2.154% | 2.273% | 0.7628% | 2.682% | 2.503% | 7.404% | 3.344% | 2.609% |  | 1.985% | 5.186% | 2.506% | 4.306% | 33.71% | 3.181%
| PDM3 |  | 4.977% | -6.009% | 12.48% | 47.7% | 9.005% | -2.103% | -4.422% | -0.8824% | 0.7044% | -6.986% | -8.573% | 1.925% | 13.16% | 12.69% |  | 8.639% | -23.34% | -4.544% | -8.939% | -18.22% | -6.009%
| Z-Score (Absolute) |  | 0.44 | 0.52 | 1.1 | NA | 0.8 | 0.18 | 0.38 | 0.07 | 0.07 | 0.6 | 0.74 | 0.18 | 1.16 | 1.12 | 23.35 | 0.77 | 2.04 | 0.39 | 0.78 | NA | NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 906b (Certified Value 0.79 ppm) | [26 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.24 | 2.25 | 2.27 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [13 empty] | 2
3
| 1 | 1 | 0.83 | 0.9 | 0.76 | 0.78 | 0.82 | 0.76446 | 0.8 | 0.63441 | 0.82 | 0.63 | 0.7 | 0.8 | [13 empty] | 1
| 1 | 2 | 0.79 | 1 | 0.76 | 0.81 | 0.8 | 0.80692 | 0.8 | 0.60787 | 0.78 | 0.6 | 0.8 | 0.8 | [13 empty] | 37
| 1 | 3 | 0.79 | 0.9 | 0.78 | 0.8 | 0.83 | 0.81308 | 0.8 | 0.61669 | 0.79 | 0.63 | 0.8 | 0.7 | [13 empty] | 16
| 1 | 4 | 0.79 | 1 | 0.74 | 0.8 | 0.88 | 0.8002 | 0.9 | 0.59371 | 0.8 | 0.66 | 0.8 | 0.8 | [13 empty] | 0.79156
| 1 | 5 | 0.78 | 1 | 0.78 | 0.76 | 0.81 | 0.76283 | 0.9 | 0.60043 | 0.82 | 0.6 | 0.7 | 0.8 | [13 empty] | 65
| 1 | 6 | 0.8 | 1 | 0.75 | 0.77 | 0.83 | 0.7907 | 0.9 | 0.60005 | 0.81 | 0.7 | 0.7 | 0.8
| Mean |  | 0.79667 | 0.96667 | 0.76167 | 0.78667 | 0.82833 | 0.7897 | 0.85 | 0.60886 | 0.80333 | 0.63667 | 0.75 | 0.78333
| Median |  | 0.79 | 1 | 0.76 | 0.79 | 0.825 | 0.79545 | 0.85 | 0.60415 | 0.805 | 0.63 | 0.75 | 0.8
| Std Dev. |  | 0.017512 | 0.05164 | 0.016021 | 0.019664 | 0.027869 | 0.021511 | 0.054772 | 0.014786 | 0.01633 | 0.038297 | 0.054772 | 0.040825
| Rel.Std.Dev. |  | 2.198% | 5.342% | 2.103% | 2.5% | 3.364% | 2.724% | 6.444% | 2.428% | 2.033% | 6.015% | 7.303% | 5.212%
| PDM3 |  | 0.6451% | 22.12% | -3.777% | -0.6183% | 4.646% | -0.2354% | 7.383% | -23.08% | 1.487% | -19.57% | -5.25% | -1.039%
| Z-Score (Absolute) |  | 0.27 | NA | 0.86 | 0.05 | 1.3 | 0.05 | NA | 5.8 | 0.49 | 4.9 | NA | NA
Results from laboratories 2.04, 2.17, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 906b (Certified Value 0.2 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 |  | 0.2 | 0.19626 | 0.19 | < 0.5 | 0.24 | 0.17 | 0.22 | 0.22 | 0.25 | 0.18 | 1.06 | 0.23 | 0.24831 | < 10 | 0.15 | 0.33 | 0.2 | 0.23 | 0.2 | NA | 0.18 |  |  |  | 1
| 1 | 2 |  | 0.2 | 0.23742 | 0.16 | < 0.5 | 0.2 | 0.2 | 0.19 | 0.22 | 0.14 | 0.18 | 0.85 | 0.21 | 0.15924 | < 10 | 0.12 | 0.38 | 0.3 | 0.22 | 0.2 | NA | 0.21 |  |  |  | 11
| 1 | 3 |  | 0.2 | 0.21637 | 0.17 | < 0.5 | 0.25 | 0.18 | 0.21 | 0.23 | 0.2 | 0.18 | 0.72 | 0.2 | 0.23651 | < 10 | 0.14 | 0.37 | 0.1 | 0.23 | 0.2 | NA | 0.17 |  |  |  | 16
| 1 | 4 | 0.2 | 0.2 | 0.22461 | 0.18 | < 0.5 | 0.22 | 0.19 | 0.2 | 0.18 | 0.18 | 0.2 | 0.804 | 0.2 | 0.23755 | < 10 | 0.13 | 0.34 | 0.1 | 0.21 | 0.2 | NA | 0.21 |  |  |  | 0.20005
| 1 | 5 | 0.2 | 0.2 | 0.16539 | 0.22 | < 0.5 | 0.21 | 0.18 | 0.18 | 0.19 | 0.2 | 0.2 | 0.82 | 0.22 | 0.22664 | < 10 | 0.17 | 0.33 | 0.2 | 0.2 | 0.2 | NA | 0.21 |  |  |  | 66
| 1 | 6 |  | 0.2 | 0.17052 | 0.21 | < 0.5 | 0.25 | 0.18 | 0.19 | 0.2 | 0.17 | 0.19 | 1.03 | 0.25 | < 0.2 | < 10 | 0.16 | 0.35 | 0.2 | 0.22 | 0.2 | NA | 0.2
| Mean |  | 0.2 | 0.2 | 0.20176 | 0.18833 |  | 0.22833 | 0.18333 | 0.19833 | 0.20667 | 0.19 | 0.18833 | 0.88067 | 0.21833 | 0.22165 |  | 0.145 | 0.35 | 0.18333 | 0.21833 | 0.2 |  | 0.19667
| Median |  | 0.2 | 0.2 | 0.20632 | 0.185 |  | 0.23 | 0.18 | 0.195 | 0.21 | 0.19 | 0.185 | 0.835 | 0.215 | 0.23651 |  | 0.145 | 0.345 | 0.2 | 0.22 | 0.2 |  | 0.205
| Std Dev. |  | 0 | 3.0405e-17 | 0.029449 | 0.023166 |  | 0.02137 | 0.010328 | 0.01472 | 0.019664 | 0.036878 | 0.0098319 | 0.13474 | 0.019408 | 0.035723 |  | 0.018708 | 0.020976 | 0.075277 | 0.01169 | 3.0405e-17 |  | 0.017512
| Rel.Std.Dev. |  | 0% | 1.52e-14% | 14.6% | 12.3% |  | 9.359% | 5.633% | 7.422% | 9.515% | 19.41% | 5.22% | 15.3% | 8.889% | 16.12% |  | 12.9% | 5.993% | 41.06% | 5.354% | 1.52e-14% |  | 8.904%
| PDM3 |  | -0.02622% | -0.02622% | 0.8547% | -5.858% |  | 14.14% | -8.357% | -0.8593% | 3.306% | -5.025% | -5.858% | 340.2% | 9.138% | 10.8% |  | -27.52% | 74.95% | -8.357% | 9.138% | -0.02622% |  | -1.692%
| Z-Score (Absolute) |  | NA | NA | 0 | 0.55 | 1.96 | 1.08 | 0.75 | 0.14 | 0.2 | 0.48 | 0.55 | 27.63 | 0.67 | 0.02 | 195.25 | 2.31 | 6.03 | NA | 0.67 | NA | NA | 0.21
Results from laboratories 2.01, 2.02, 2.23 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 906b (Certified Value 14.8 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 15.11 | 14.2 | 15.1903 | 14.4 | 13.38 | 16.1 | 14.8 | 15.4 | 15.25 | 16.25 | 14.13 | 14.9 | 14.8 | 14.8906 | 14.31 | 10.0615 | 12 | 15.7 | NA | 9.2 | 12 | 14.8 |  |  |  | 1
| 1 | 2 | 15.27 | 14.4 | 15.9798 | 15.8 | 13.6 | 14.6 | 14.45 | 15.35 | 15.5 | 14.9 | 14.44 | 7.7 | 14.5 | 15.5551 | 14.55 | 10.1129 | 12 | 15.5 | NA | 8.2 | 15.2 | 15.2 |  |  |  | 39
| 1 | 3 | 14.56 | 14.2 | 14.9896 | 14.5 | 13.89 | 15.85 | 15.05 | 15.7 | 15.45 | 14.65 | 14.46 | 13.9 | 14.5 | 16.136 | 14.15 | 10.1501 | 13 | 15.3 | NA | 8.5 | 14.9 | 14.3 |  |  |  | 16
| 1 | 4 | 14.83 | 14.6 | 15.2625 | 15.4 | 13.47 | 14.75 | 15.05 | 15.4 | 15.25 | 14.2 | 14.53 | 13.7 | 15 | 15.8622 | 14.53 | 10.1879 | 14 | 15.4 | NA | 9.1 | 13.4 | 14.9 |  |  |  | 14.8275
| 1 | 5 | 14.64 | 14.6 | 15.8155 | 15.3 | 14.02 | 14.85 | 14.4 | 15.8 | 15.7 | 15.25 | 13.94 | 14.2 | 14 | 15.1419 | 14.59 | 10.2572 | 14 | 15.4 | NA | 8.5 | 11.6 | 15 |  |  |  | 67
| 1 | 6 | 14.98 | 14.6 | 15.507 | 16 | 13.36 | 14.9 | 14 | 15 | 14.95 | 16.2 | 14.46 | 14.4 | 14.7 | 15.6324 | 14.52 | 9.9662 | 12 | 14.9 | NA | 10 | 12.7 | 14.7
| Mean |  | 14.8983 | 14.4333 | 15.4575 | 15.2333 | 13.62 | 15.175 | 14.625 | 15.4417 | 15.35 | 15.2417 | 14.3267 | 13.1333 | 14.5833 | 15.5364 | 14.4417 | 10.1226 | 12.8333 | 15.3667 |  | 8.9167 | 13.3 | 14.8167
| Median |  | 14.905 | 14.5 | 15.3848 | 15.35 | 13.535 | 14.875 | 14.625 | 15.4 | 15.35 | 15.075 | 14.45 | 14.05 | 14.6 | 15.5938 | 14.525 | 10.1315 | 12.5 | 15.4 |  | 8.8 | 13.05 | 14.85
| Std Dev. |  | 0.27404 | 0.19664 | 0.38256 | 0.65929 | 0.27604 | 0.63305 | 0.41563 | 0.28358 | 0.25884 | 0.83511 | 0.23577 | 2.6942 | 0.34303 | 0.45781 | 0.17325 | 0.10142 | 0.98319 | 0.26583 |  | 0.65549 | 1.4913 | 0.30605
| Rel.Std.Dev. |  | 1.839% | 1.362% | 2.475% | 4.328% | 2.027% | 4.172% | 2.842% | 1.836% | 1.686% | 5.479% | 1.646% | 20.51% | 2.352% | 2.947% | 1.2% | 1.002% | 7.661% | 1.73% |  | 7.351% | 11.21% | 2.066%
| PDM3 |  | 0.4779% | -2.658% | 4.249% | 2.737% | -8.143% | 2.344% | -1.366% | 4.142% | 3.524% | 2.793% | -3.378% | -11.43% | -1.647% | 4.781% | -2.602% | -31.73% | -13.45% | 3.636% |  | -39.86% | -10.3% | -0.07288%
| Z-Score (Absolute) |  | 0.21 | 0.34 | 0.86 | 0.6 | 1.29 | 0.53 | 0.11 | 0.85 | 0.74 | 0.61 | 0.46 | 1.86 | 0.16 | 0.96 | 0.33 | 5.39 | NA | 0.76 | NA | 6.81 | 1.67 | 0.11
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 906b (Certified Value 0.111 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.1122 | 0.104 | 0.11329 | 0.114 | 0.15 | 0.11 | 0.114 | 0.113 | 0.118 | 0.115 | 0.111 | 0.124 | 0.11 | 0.10845 | 0.104 | 0.11 | 0.11269 | 0.11 | 0.119 | 0.1 | 0.1156 | 0.1 | 0.116 | 0.115 |  | 1
| 1 | 2 | 0.1116 | 0.101 | 0.11133 | 0.114 | 0.14 | 0.111 | 0.113 | 0.118 | 0.119 | 0.114 | 0.108 | 0.124 | 0.1 | 0.10682 | 0.105 | 0.11 | 0.11189 | 0.11 | 0.107 | 0.1 | 0.114 | 0.11 | 0.109 | 0.109 |  | 40
| 1 | 3 | 0.1103 | 0.103 | 0.11393 | 0.112 | 0.14 | 0.112 | 0.112 | 0.119 | 0.118 | 0.111 | 0.11 | 0.118 | 0.1 | 0.10649 | 0.101 | 0.11 | 0.11213 | 0.11 | 0.116 | 0.09 | 0.1132 | 0.1 | 0.105 | 0.108 |  | 16
| 1 | 4 | 0.1088 | 0.106 | 0.11437 | 0.114 | 0.13 | 0.11 | 0.112 | 0.119 | 0.119 | 0.111 | 0.108 | 0.119 | 0.11 | 0.1081 | 0.103 | 0.11 | 0.1138 | 0.11 | 0.108 | 0.09 | 0.1154 | 0.1 | 0.101 | 0.11 |  | 0.1106
| 1 | 5 | 0.1109 | 0.1 | 0.11182 | 0.113 | 0.13 | 0.113 | 0.115 | 0.115 | 0.12 | 0.115 | 0.109 | 0.124 | 0.11 | 0.1081 | 0.102 | 0.11 | 0.11324 | 0.11 | 0.114 | 0.1 | 0.1168 | 0.1 | 0.116 | 0.111 |  | 68
| 1 | 6 | 0.1106 | 0.106 | 0.11072 | 0.116 | 0.14 | 0.109 | 0.111 | 0.119 | 0.116 | 0.115 | 0.11 | 0.122 | 0.11 | 0.10821 | 0.105 | 0.11 | 0.11143 | 0.11 | 0.109 | 0.1 | 0.1136 | 0.1 | 0.116 | 0.112
| Mean |  | 0.11073 | 0.10333 | 0.11258 | 0.11383 | 0.13833 | 0.11083 | 0.11283 | 0.11717 | 0.11833 | 0.1135 | 0.10933 | 0.12183 | 0.10667 | 0.10769 | 0.10333 | 0.11 | 0.11253 | 0.11 | 0.11217 | 0.096667 | 0.11477 | 0.10167 | 0.1105 | 0.11083
| Median |  | 0.11075 | 0.1035 | 0.11255 | 0.114 | 0.14 | 0.1105 | 0.1125 | 0.1185 | 0.1185 | 0.1145 | 0.1095 | 0.123 | 0.11 | 0.1081 | 0.1035 | 0.11 | 0.11241 | 0.11 | 0.1115 | 0.1 | 0.1147 | 0.1 | 0.1125 | 0.1105
| Std Dev. |  | 0.0011725 | 0.0025033 | 0.0014905 | 0.0013292 | 0.0075277 | 0.001472 | 0.001472 | 0.0025626 | 0.0013663 | 0.0019748 | 0.0012111 | 0.0027142 | 0.005164 | 0.00081934 | 0.001633 | 0 | 0.0008853 | 0 | 0.0048751 | 0.005164 | 0.001388 | 0.0040825 | 0.0065345 | 0.0024833
| Rel.Std.Dev. |  | 1.059% | 2.423% | 1.324% | 1.168% | 5.442% | 1.328% | 1.305% | 2.187% | 1.155% | 1.74% | 1.108% | 2.228% | 4.841% | 0.7608% | 1.58% | 0% | 0.7867% | 0% | 4.346% | 5.342% | 1.209% | 4.016% | 5.914% | 2.241%
| PDM3 |  | 0.1191% | -6.572% | 1.784% | 2.922% | 25.07% | 0.2095% | 2.018% | 5.936% | 6.991% | 2.621% | -1.147% | 10.16% | -3.558% | -2.629% | -6.572% | -0.544% | 1.745% | -0.544% | 1.415% | -12.6% | 3.766% | -8.079% | -0.09188% | 0.2095%
| Z-Score (Absolute) |  | 0.02 | 1.78 | 0.41 | 0.71 | 6.54 | 0 | 0.48 | 1.51 | 1.78 | 0.63 | 0.36 | 2.62 | 0.99 | 0.75 | 1.78 | 0.2 | 0.4 | 0.2 | 0.32 | 3.37 | 0.94 | 2.18 | 0.08 | 0
| Analytical results for Tl in OREAS 906b (Certified Value 0.72 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.77 | 0.7 | 0.75677 | 0.7 | 0.7 | 0.71 | 0.66 | 0.7 | 0.76 | 0.72 | 0.7 | 0.84 | 0.75 | 0.69082 | 0.75 | < 5 | 0.72 | 0.7 | 0.7 | 0.7 | 0.77 | 0.75 |  |  |  | 1
| 1 | 2 | 0.76 | 0.7 | 0.73795 | 0.79 | 0.66 | 0.7 | 0.66 | 0.68 | 0.75 | 0.73 | 0.69 | 0.74 | 0.75 | 0.70617 | 0.74 | < 5 | 0.71 | 0.75 | 0.72 | 0.69 | 0.758 | 0.75 |  |  |  | 41
| 1 | 3 | 0.74 | 0.7 | 0.73867 | 0.71 | 0.7 | 0.72 | 0.69 | 0.7 | 0.78 | 0.72 | 0.71 | 0.74 | 0.74 | 0.71366 | 0.73 | < 5 | 0.69 | 0.78 | 0.69 | 0.7 | 0.749 | 0.72 |  |  |  | 16
| 1 | 4 | 0.75 | 0.7 | 0.7192 | 0.75 | 0.7 | 0.69 | 0.66 | 0.69 | 0.77 | 0.61 | 0.69 | 0.77 | 0.71 | 0.72225 | 0.79 | < 5 | 0.71 | 0.73 | 0.69 | 0.71 | 0.772 | 0.75 |  |  |  | 0.7211
| 1 | 5 | 0.72 | 0.73 | 0.71139 | 0.77 | 0.72 | 0.71 | 0.67 | 0.72 | 0.77 | 0.72 | 0.68 | 0.75 | 0.74 | 0.69499 | 0.73 | < 5 | 0.71 | 0.7 | 0.69 | 0.72 | 0.765 | 0.73 |  |  |  | 69
| 1 | 6 | 0.77 | 0.71 | 0.7599 | 0.78 | 0.64 | 0.69 | 0.65 | 0.68 | 0.74 | 0.75 | 0.69 | 0.79 | 0.72 | 0.71164 | 0.73 | < 5 | 0.7 | 0.7 | 0.71 | 0.72 | 0.747 | 0.73
| Mean |  | 0.75167 | 0.70667 | 0.73731 | 0.75 | 0.68667 | 0.70333 | 0.665 | 0.695 | 0.76167 | 0.70833 | 0.69333 | 0.77167 | 0.735 | 0.70659 | 0.745 |  | 0.70667 | 0.72667 | 0.7 | 0.70667 | 0.76017 | 0.73833
| Median |  | 0.755 | 0.7 | 0.73831 | 0.76 | 0.7 | 0.705 | 0.66 | 0.695 | 0.765 | 0.72 | 0.69 | 0.76 | 0.74 | 0.7089 | 0.735 |  | 0.71 | 0.715 | 0.695 | 0.705 | 0.7615 | 0.74
| Std Dev. |  | 0.019408 | 0.012111 | 0.019447 | 0.037417 | 0.030111 | 0.012111 | 0.013784 | 0.015166 | 0.01472 | 0.049565 | 0.010328 | 0.038687 | 0.016432 | 0.011866 | 0.023452 |  | 0.010328 | 0.033267 | 0.012649 | 0.012111 | 0.01061 | 0.013292
| Rel.Std.Dev. |  | 2.582% | 1.714% | 2.638% | 4.989% | 4.385% | 1.722% | 2.073% | 2.182% | 1.933% | 6.997% | 1.49% | 5.013% | 2.236% | 1.679% | 3.148% |  | 1.462% | 4.578% | 1.807% | 1.714% | 1.396% | 1.8%
| PDM3 |  | 4.239% | -2.001% | 2.249% | 4.008% | -4.775% | -2.464% | -7.78% | -3.619% | 5.626% | -1.77% | -3.85% | 7.013% | 1.928% | -2.012% | 3.315% |  | -2.001% | 0.7722% | -2.926% | -2.001% | 5.418% | 2.39%
| Z-Score (Absolute) |  | 1.06 | 0.34 | 0.62 | 1.01 | 0.96 | 0.44 | 1.63 | 0.7 | 1.37 | 0.29 | 0.75 | 1.69 | 0.54 | 0.34 | 0.86 | 55.47 | 0.34 | 0.29 | 0.54 | 0.34 | 1.33 | 0.65
| Analytical results for Tm in OREAS 906b (Certified Value 0.11 ppm) | [26 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.04 | 2.15 | 2.18 | 2.20 | 2.25 | 2.27 | [18 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [18 empty] | 2
2
| 1 | 1 | 0.12 | 0.1 | 0.12417 | < 0.1 | 0.12 | < 0.1 | 0.1 | [18 empty] | 1
| 1 | 2 | 0.12 | 0.1 | 0.11971 | 0.10525 | 0.11 | 0.1 | 0.1 | [18 empty] | 19
| 1 | 3 | 0.13 | 0.1 | 0.12833 | < 0.1 | 0.12 | 0.1 | 0.1 | [18 empty] | 16
| 1 | 4 | 0.12 | 0.1 | 0.12262 | < 0.1 | 0.11 | 0.1 | 0.1 | [18 empty] | 0.10679
| 1 | 5 | 0.12 | 0.2 | 0.12004 | < 0.1 | 0.11 | 0.1 | 0.1 | [18 empty] | 70
| 1 | 6 | 0.13 | 0.1 | 0.11973 | < 0.1 | 0.1 | < 0.1 | 0.1
| Mean |  | 0.12333 | 0.11667 | 0.12243 | 0.10525 | 0.11167 | 0.1 | 0.1
| Median |  | 0.12 | 0.1 | 0.12133 | 0.10525 | 0.11 | 0.1 | 0.1
| Std Dev. |  | 0.005164 | 0.040825 | 0.0034121 |  | 0.0075277 | 0 | 1.5202e-17
| Rel.Std.Dev. |  | 4.187% | 34.99% | 2.787% |  | 6.741% | 0% | 1.52e-14%
| PDM3 |  | 15.49% | 9.244% | 14.64% | -1.444% | 4.562% | -6.362% | -6.362%
| Z-Score (Absolute) |  | 0.67 | 0.29 | 0.62 | 3.03 | 0 | 1.64 | 0.67
| Analytical results for U in OREAS 906b (Certified Value 5.08 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.26 | 5.08 | 5.0785 | 5 | 4.76 | 5 | 4.9 | 5.5 | 5.1 | 5 | 5.07 | 5.38 | 5.3 | 5.3071 | 5.04 | 3.4116 | < 10 | 5.32 | 5.06 | 4.01 | 4.5 | 5.22 | 5.3 |  |  | 1
| 1 | 2 | 5.27 | 5.09 | 5.3552 | 5.2 | 4.78 | 5 | 4.8 | 5 | 5.2 | 4.8 | 5.09 | 5.02 | 5.3 | 5.6005 | 4.93 | 3.4352 | < 10 | 5.03 | 5.07 | 3.45 | 5.1 | 5.14 | 5.4 |  |  | 43
| 1 | 3 | 5.1 | 5 | 4.9755 | 4.9 | 4.79 | 5 | 5.1 | 5.4 | 5.2 | 4.7 | 5.18 | 4.98 | 5.4 | 5.7867 | 5.01 | 3.4305 | < 10 | 5.08 | 5.12 | 3.2 | 5.1 | 5.13 | 5 |  |  | 16
| 1 | 4 | 5.17 | 5.07 | 5.0514 | 5 | 4.75 | 4.8 | 4.9 | 5 | 5.1 | 4.8 | 5.07 | 5.02 | 5.2 | 5.6249 | 4.96 | 3.4898 | < 10 | 5.17 | 5.14 | 3.8 | 4.8 | 5.26 | 5.2 |  |  | 5.0827
| 1 | 5 | 5.21 | 5.07 | 5.2627 | 5.2 | 4.75 | 5 | 4.8 | 5.2 | 5.2 | 4.9 | 4.9 | 5.16 | 5.3 | 5.3928 | 5.04 | 3.4747 | < 10 | 5.14 | 5.03 | 3.72 | 4.5 | 5.2 | 5.4 |  |  | 71
| 1 | 6 | 5.33 | 5.12 | 5.2389 | 5.3 | 4.77 | 5 | 4.6 | 4.8 | 5 | 4.9 | 5.11 | 5.19 | 5.2 | 5.5106 | 4.89 | 3.4558 | < 10 | 4.93 | 5 | 4.16 | 4.8 | 5.09 | 5.2
| Mean |  | 5.2233 | 5.0717 | 5.1604 | 5.1 | 4.7667 | 4.9667 | 4.85 | 5.15 | 5.1333 | 4.85 | 5.07 | 5.125 | 5.2833 | 5.5371 | 4.9783 | 3.4496 |  | 5.1117 | 5.07 | 3.7233 | 4.8 | 5.1733 | 5.25
| Median |  | 5.235 | 5.075 | 5.1587 | 5.1 | 4.765 | 5 | 4.85 | 5.1 | 5.15 | 4.85 | 5.08 | 5.09 | 5.3 | 5.5555 | 4.985 | 3.4455 |  | 5.11 | 5.065 | 3.76 | 4.8 | 5.17 | 5.25
| Std Dev. |  | 0.081404 | 0.039707 | 0.14651 | 0.15492 | 0.01633 | 0.08165 | 0.16432 | 0.26646 | 0.08165 | 0.10488 | 0.092736 | 0.1507 | 0.075277 | 0.17233 | 0.061779 | 0.029336 |  | 0.13288 | 0.052915 | 0.35421 | 0.26833 | 0.06377 | 0.15166
| Rel.Std.Dev. |  | 1.558% | 0.7829% | 2.839% | 3.038% | 0.3426% | 1.644% | 3.388% | 5.174% | 1.591% | 2.162% | 1.829% | 2.94% | 1.425% | 3.112% | 1.241% | 0.8504% |  | 2.6% | 1.044% | 9.513% | 5.59% | 1.233% | 2.889%
| PDM3 |  | 2.766% | -0.2179% | 1.527% | 0.3395% | -6.219% | -2.284% | -4.579% | 1.323% | 0.9953% | -4.579% | -0.2507% | 0.8314% | 3.946% | 8.939% | -2.054% | -32.13% |  | 0.569% | -0.2507% | -26.75% | -5.563% | 1.782% | 3.291%
| Z-Score (Absolute) |  | 1.01 | 0 | 0.59 | 0.19 | 2.02 | 0.7 | 1.47 | 0.52 | 0.41 | 1.47 | 0.01 | 0.35 | 1.4 | 3.09 | 0.62 | 10.76 | 0.48 | 0.27 | 0.01 | 8.94 | 1.8 | 0.67 | 1.18
| Analytical results for V in OREAS 906b (Certified Value 3.88 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 4 | 4 | 5.326 | 5 | 6.97 | 4 | 4 | 4 | 3 | 4 | 4 | 8 | 5 | 3.7682 | < 2 | 3.7 | 3 | 3 | 4 | 5.56 | 5 | 5 | 4 | 4 |  | 1
| 1 | 2 | 4 | 4 | 6.0242 | 4 | 7.15 | 4 | 4 | 3 | 4 | 4 | 4 | 8 | 6 | 3.8351 | < 2 | 3.8 | 3 | 3 | 4 | 5 | 6 | 5 | 4 | 4 |  | 12
| 1 | 3 | 4 | 4 | 6.0615 | 5 | 7.37 | 4 | 4 | 3 | 4 | 4 | 4 | 8 | 5 | 3.7873 | < 2 | 3.7 | 3 | 3 | 4 | 5.78 | 5 | 5 | 4 | 4 |  | 16
| 1 | 4 | 4 | 4 | 6.0914 | 5 | 7.05 | 4 | 4 | 3 | 4 | 3 | 4 | 8.1 | 5 | 3.8141 | < 2 | 4.1 | 4 | 3 | 4 | 5.87 | 5 | 6 | 4 | 4 |  | 3.8836
| 1 | 5 | 3 | 4 | 5.258 | 5 | 7.34 | 4 | 4 | 4 | 3 | 4 | 4 | 8 | 6 | 3.7708 | < 2 | 3.8 | 3 | 3 | 4 | 5.67 | 5 | 6 | 4 | 4 |  | 72
| 1 | 6 | 3 | 4 | 5.2378 | 5 | 7.13 | 4 | 4 | 3 | 4 | 6 | 4 | 8 | 5 | 3.9483 | < 2 | 3.8 | 3 | 3 | 4 | 5.91 | 5 | 5 | 4 | 4
| Mean |  | 3.6667 | 4 | 5.6665 | 4.8333 | 7.1683 | 4 | 4 | 3.3333 | 3.6667 | 4.1667 | 4 | 8.0167 | 5.3333 | 3.8206 |  | 3.8167 | 3.1667 | 3 | 4 | 5.6317 | 5.1667 | 5.3333 | 4 | 4
| Median |  | 4 | 4 | 5.6751 | 5 | 7.14 | 4 | 4 | 3 | 4 | 4 | 4 | 8 | 5 | 3.8007 |  | 3.8 | 3 | 3 | 4 | 5.725 | 5 | 5 | 4 | 4
| Std Dev. |  | 0.5164 | 0 | 0.43155 | 0.40825 | 0.15829 | 0 | 0 | 0.5164 | 0.5164 | 0.98319 | 0 | 0.040825 | 0.5164 | 0.067673 |  | 0.1472 | 0.40825 | 0 | 0 | 0.33523 | 0.40825 | 0.5164 | 0 | 0
| Rel.Std.Dev. |  | 14.08% | 0% | 7.616% | 8.447% | 2.208% | 0% | 0% | 15.49% | 14.08% | 23.6% | 0% | 0.5092% | 9.682% | 1.771% |  | 3.857% | 12.89% | 0% | 0% | 5.953% | 7.902% | 9.682% | 0% | 0%
| PDM3 |  | -5.587% | 2.996% | 45.91% | 24.45% | 84.58% | 2.996% | 2.996% | -14.17% | -5.587% | 7.288% | 2.996% | 106.4% | 37.33% | -1.623% |  | -1.724% | -18.46% | -22.75% | 2.996% | 45.01% | 33.04% | 37.33% | 2.996% | 2.996%
| Z-Score (Absolute) |  | 0.45 | 0 | 2.25 | 1.12 | 4.27 | 0 | 0 | 0.9 | 0.45 | 0.22 | 0 | 5.42 | 1.8 | 0.24 | 4.05 | 0.25 | 1.12 | 1.35 | 0 | 2.2 | 1.57 | 1.8 | 0 | 0
| Analytical results for W in OREAS 906b (Certified Value 2.61 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.6 | 2.6 | 2.7389 | 2.4 | 2.11 | 2.6 | 2.4 | 2.5 | 2.9 | 2.4 | 2.5 | 3.66 | 2.6 | 2.4782 | < 10 | 2.7 | < 20 | 2.8 | 10.2 | 2.67 | 2 | 2.6 | 2.6 |  |  | 1
| 1 | 2 | 2.6 | 2.5 | 2.6894 | 2.5 | 1.93 | 2.6 | 2.4 | 2.4 | 2.9 | 2.5 | 2.5 | 3.32 | 2.5 | 2.8546 | < 10 | 2.8 | < 20 | 2.8 | 10.7 | 2.72 | 3 | 2.4 | 2.4 |  |  | 13
| 1 | 3 | 2.7 | 2.5 | 2.8492 | 2.4 | 2.83 | 2.8 | 2.4 | 2.4 | 2.8 | 2.5 | 2.6 | 3.38 | 2.5 | 2.8947 | < 10 | 2.8 | < 20 | 2.6 | 10.7 | 2.8 | 2 | 2.7 | 2.4 |  |  | 16
| 1 | 4 | 2.5 | 2.6 | 2.6589 | 2.6 | 2.17 | 2.7 | 2.4 | 2.4 | 2.8 | 2.4 | 2.5 | 3.24 | 2.8 | 2.8912 | < 10 | 2.9 | < 20 | 2.7 | 10.6 | 2.66 | 3 | 2.8 | 2.5 |  |  | 2.6107
| 1 | 5 | 2.6 | 2.5 | 2.8116 | 2.5 | 1.96 | 2.6 | 2.4 | 2.5 | 2.8 | 2.4 | 2.4 | 3.08 | 2.5 | 2.4571 | < 10 | 2.8 | < 20 | 2.6 | 10.9 | 2.79 | 3 | 2.9 | 2.6 |  |  | 73
| 1 | 6 | 2.6 | 2.6 | 2.8458 | 2.6 | 2.1 | 2.7 | 2.3 | 2.4 | 2.7 | 2.5 | 2.5 | 3.36 | 2.6 | 2.7788 | < 10 | 2.8 | < 20 | 2.7 | 11.2 | 2.6 | 2 | 2.8 | 2.5
| Mean |  | 2.6 | 2.55 | 2.7656 | 2.5 | 2.1833 | 2.6667 | 2.3833 | 2.4333 | 2.8167 | 2.45 | 2.5 | 3.34 | 2.5833 | 2.7258 |  | 2.8 |  | 2.7 | 10.7167 | 2.7067 | 2.5 | 2.7 | 2.5
| Median |  | 2.6 | 2.55 | 2.7752 | 2.5 | 2.105 | 2.65 | 2.4 | 2.4 | 2.8 | 2.45 | 2.5 | 3.34 | 2.55 | 2.8167 |  | 2.8 |  | 2.7 | 10.7 | 2.695 | 2.5 | 2.75 | 2.5
| Std Dev. |  | 0.063246 | 0.054772 | 0.081791 | 0.089443 | 0.33007 | 0.08165 | 0.040825 | 0.05164 | 0.075277 | 0.054772 | 0.063246 | 0.191 | 0.1169 | 0.20437 |  | 0.063246 |  | 0.089443 | 0.33116 | 0.078401 | 0.54772 | 0.17889 | 0.089443
| Rel.Std.Dev. |  | 2.433% | 2.148% | 2.957% | 3.578% | 15.12% | 3.062% | 1.713% | 2.122% | 2.673% | 2.236% | 2.53% | 5.718% | 4.525% | 7.498% |  | 2.259% |  | 3.313% | 3.09% | 2.897% | 21.91% | 6.625% | 3.578%
| PDM3 |  | -0.4087% | -2.324% | 5.936% | -4.239% | -16.37% | 2.145% | -8.708% | -6.793% | 7.891% | -6.154% | -4.239% | 27.94% | -1.047% | 4.408% |  | 7.252% |  | 3.422% | 310.5% | 3.677% | -4.239% | 3.422% | -4.239%
| Z-Score (Absolute) |  | 0.35 | 0.57 | 0.35 | 0.78 | 2.13 | 0.07 | 1.28 | 1.06 | 0.57 | 0.99 | 0.78 | 2.8 | 0.43 | 0.18 | 9.87 | 0.5 | 31.16 | 0.07 | 34.21 | 0.1 | NA | 0.07 | 0.78
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 906b (Certified Value 15.8 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 15.73 | 15.5 | 16.1965 | 15.2 | 14.3 | 15.9 | 16 | 16.4 | 15.8 | 15.2 | 15.5 | 17.5 | 15.7 | 16.2621 | 15.9 | 14.9 | 11.8616 | 15 | 15.9 | 15.9 | 16.2 | 9.7 | 13.9 | 16.8 |  | 1
| 1 | 2 | 15.38 | 15.7 | 16.3794 | 15.8 | 14.1 | 15.7 | 15 | 16.3 | 16.2 | 14.3 | 16 | 13 | 15.7 | 16.2961 | 16 | 14.9 | 11.8723 | 16 | 15.3 | 17.2 | 15.8 | 9.1 | 16.6 | 15.9 |  | 46
| 1 | 3 | 15.33 | 15.8 | 15.9003 | 15.7 | 14.4 | 16 | 15.6 | 16.7 | 16.2 | 15.8 | 15.7 | 16.6 | 15.5 | 16.2871 | 15.8 | 14.9 | 11.871 | 15 | 15.7 | 15.6 | 16.4 | 9.7 | 16.6 | 15.7 |  | 16
| 1 | 4 | 15.53 | 15.9 | 16.0185 | 15.8 | 13.9 | 15.7 | 15.2 | 16.7 | 16.1 | 14.3 | 15.3 | 16.5 | 15.9 | 16.1221 | 15.7 | 15.4 | 11.8469 | 16 | 16.3 | 16.4 | 15.2 | 10 | 15.1 | 15.5 |  | 15.7894
| 1 | 5 | 15.26 | 15.9 | 16.1392 | 15.6 | 14.1 | 15.9 | 15.4 | 16.4 | 16.2 | 14.8 | 15.3 | 17 | 15.1 | 16.0997 | 15.7 | 15.3 | 11.896 | 16 | 16.1 | 16.6 | 15.3 | 9.2 | 12.8 | 16.1 |  | 74
| 1 | 6 | 15.81 | 15.6 | 16.2876 | 15.6 | 14.4 | 15.9 | 14.8 | 16.7 | 15.5 | 14.7 | 15.2 | 17.1 | 16.1 | 16.4791 | 16.2 | 15 | 11.8843 | 15 | 15.5 | 14.2 | 15.8 | 10.9 | 13.7 | 16.3
| Mean |  | 15.5067 | 15.7333 | 16.1536 | 15.6167 | 14.2 | 15.85 | 15.3333 | 16.5333 | 16 | 14.85 | 15.5 | 16.2833 | 15.6667 | 16.2577 | 15.8833 | 15.0667 | 11.872 | 15.5 | 15.8 | 15.9833 | 15.7833 | 9.7667 | 14.7833 | 16.05
| Median |  | 15.455 | 15.75 | 16.1678 | 15.65 | 14.2 | 15.9 | 15.3 | 16.55 | 16.15 | 14.75 | 15.4 | 16.8 | 15.7 | 16.2746 | 15.85 | 14.95 | 11.8716 | 15.5 | 15.8 | 16.15 | 15.8 | 9.7 | 14.5 | 16
| Std Dev. |  | 0.22385 | 0.1633 | 0.17515 | 0.22286 | 0.2 | 0.12247 | 0.43205 | 0.18619 | 0.28983 | 0.57533 | 0.30332 | 1.6485 | 0.34448 | 0.13763 | 0.19408 | 0.22509 | 0.017107 | 0.54772 | 0.37417 | 1.0362 | 0.47504 | 0.65013 | 1.5867 | 0.46368
| Rel.Std.Dev. |  | 1.444% | 1.038% | 1.084% | 1.427% | 1.408% | 0.7727% | 2.818% | 1.126% | 1.811% | 3.874% | 1.957% | 10.12% | 2.199% | 0.8466% | 1.222% | 1.494% | 0.1441% | 3.534% | 2.368% | 6.483% | 3.01% | 6.657% | 10.73% | 2.889%
| PDM3 |  | -1.791% | -0.3551% | 2.306% | -1.094% | -10.07% | 0.3838% | -2.888% | 4.712% | 1.334% | -5.95% | -1.833% | 3.128% | -0.7773% | 2.966% | 0.5949% | -4.577% | -24.81% | -1.833% | 0.06715% | 1.228% | -0.0384% | -38.14% | -6.372% | 1.65%
| Z-Score (Absolute) |  | 0.48 | 0 | 0.89 | 0.25 | 3.27 | 0.25 | 0.85 | 1.7 | 0.57 | 1.88 | 0.5 | 1.17 | 0.14 | 1.12 | 0.32 | 1.42 | 8.22 | NA | 0.14 | 0.53 | 0.11 | 12.71 | 2.02 | 0.67
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 906b (Certified Value 0.61 ppm) | [26 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.24 | 2.25 | 2.27 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [13 empty] | 2
2
| 1 | 1 | 0.66 | 0.7 | 0.56 | 0.7 | 0.6 | 0.60436 | 0.6 | 0.48749 | 0.61 | 0.6 | 0.5 | 0.6 | [13 empty] | 1
| 1 | 2 | 0.62 | 0.7 | 0.58 | 0.7 | 0.6 | 0.63393 | 0.6 | 0.50547 | 0.58 | 0.5 | 0.6 | 0.7 | [13 empty] | 47
| 1 | 3 | 0.66 | 0.7 | 0.56 | 0.7 | 0.7 | 0.65677 | 0.6 | 0.50673 | 0.56 | 0.5 | 0.6 | 0.6 | [13 empty] | 16
| 1 | 4 | 0.63 | 0.7 | 0.58 | 0.7 | 0.7 | 0.64279 | 0.6 | 0.52084 | 0.6 | 0.6 | 0.6 | 0.7 | [13 empty] | 0.6111
| 1 | 5 | 0.63 | 0.8 | 0.57 | 0.6 | 0.6 | 0.6375 | 0.6 | 0.5696 | 0.6 | 0.5 | 0.5 | 0.7 | [13 empty] | 75
| 1 | 6 | 0.68 | 0.7 | 0.56 | 0.6 | 0.6 | 0.65017 | 0.6 | 0.51572 | 0.59 | 0.6 | 0.6 | 0.6
| Mean |  | 0.64667 | 0.71667 | 0.56833 | 0.66667 | 0.63333 | 0.63759 | 0.6 | 0.51764 | 0.59 | 0.55 | 0.56667 | 0.65
| Median |  | 0.645 | 0.7 | 0.565 | 0.7 | 0.6 | 0.64015 | 0.6 | 0.51122 | 0.595 | 0.55 | 0.6 | 0.65
| Std Dev. |  | 0.023381 | 0.040825 | 0.0098319 | 0.05164 | 0.05164 | 0.018279 | 0 | 0.027893 | 0.017889 | 0.054772 | 0.05164 | 0.054772
| Rel.Std.Dev. |  | 3.616% | 5.696% | 1.73% | 7.746% | 8.154% | 2.867% | 0% | 5.388% | 3.032% | 9.959% | 9.113% | 8.427%
| PDM3 |  | 5.821% | 17.28% | -6.998% | 9.093% | 3.639% | 4.335% | -1.816% | -15.29% | -3.452% | -9.998% | -7.271% | 6.366%
| Z-Score (Absolute) |  | 0.5 | 1.65 | 0.8 | 0.83 | 0.28 | 0.35 | 0.28 | 1.64 | 0.44 | 1.1 | 0.83 | 0.55
| Analytical results for Zn in OREAS 906b (Certified Value 147 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
0
| 1 | 1 | 153 | 143 | 148.0823 | 146.4 | 150.29 | 148 | 147 | 153 | 154 | 147 | 145 | 169 | 142 | 149.3249 | 137 | 144.9 | 139.35 | 145 | 143 | 134 | 168 | 141 | 154 | 147 | 159.4 | 1
| 1 | 2 | 150 | 144 | 153.3159 | 147.8 | 149.91 | 148 | 147 | 157 | 156 | 144 | 145 | 157 | 145 | 150.4804 | 142 | 143.2 | 136.79 | 148 | 142 | 148 | 159 | 141 | 152 | 148 | 152.4 | 28
| 1 | 3 | 155 | 145 | 146.1866 | 148.4 | 149.7 | 149 | 146 | 158 | 156 | 145 | 145 | 154 | 139 | 150.6165 | 137 | 145.3 | 140.74 | 147 | 146 | 152 | 158 | 140 | 147 | 148 | 148 | 16
| 1 | 4 | 153 | 144 | 150.6317 | 144.4 | 151.86 | 146 | 147 | 158 | 155 | 142 | 144 | 155 | 140 | 150.5463 | 141 | 145.2 | 139.07 | 152 | 146 | 139 | 160 | 141 | 144 | 147 | 147.3 | 146.8536
| 1 | 5 | 151 | 145 | 150.9191 | 147.1 | 153 | 150 | 150 | 153 | 156 | 147 | 145 | 158 | 144 | 149.7546 | 134 | 149.3 | 141.77 | 149 | 148 | 133 | 166 | 141 | 156 | 148 | 148.4 | 76
| 1 | 6 | 154 | 143 | 154.298 | 148.8 | 150.76 | 145 | 146 | 156 | 150 | 145 | 145 | 166 | 138 | 150.8758 | 139 | 142.5 | 136.62 | 151 | 143 | 130 | 158 | 141 | 157 | 146 | 143.9
| Mean |  | 152.6667 | 144 | 150.5723 | 147.15 | 150.92 | 147.6667 | 147.1667 | 155.8333 | 154.5 | 145 | 144.8333 | 159.8333 | 141.3333 | 150.2664 | 138.3333 | 145.0667 | 139.0567 | 148.6667 | 144.6667 | 139.3333 | 161.5 | 140.8333 | 151.6667 | 147.3333 | 149.9
| Median |  | 153 | 144 | 150.7754 | 147.45 | 150.525 | 148 | 147 | 156.5 | 155.5 | 145 | 145 | 157.5 | 141 | 150.5133 | 138 | 145.05 | 139.21 | 148.5 | 144.5 | 136.5 | 159.5 | 141 | 153 | 147.5 | 148.2
| Std Dev. |  | 1.8619 | 0.89443 | 3.0648 | 1.6022 | 1.276 | 1.8619 | 1.472 | 2.3166 | 2.3452 | 1.8974 | 0.40825 | 6.1779 | 2.8048 | 0.59437 | 2.9439 | 2.3704 | 2.0667 | 2.582 | 2.3381 | 8.8468 | 4.3704 | 0.40825 | 5.164 | 0.8165 | 5.3874
| Rel.Std.Dev. |  | 1.22% | 0.6211% | 2.035% | 1.089% | 0.8455% | 1.261% | 1% | 1.487% | 1.518% | 1.309% | 0.2819% | 3.865% | 1.984% | 0.3955% | 2.128% | 1.634% | 1.486% | 1.737% | 1.616% | 6.349% | 2.706% | 0.2899% | 3.405% | 0.5542% | 3.594%
| PDM3 |  | 3.958% | -1.943% | 2.532% | 0.2019% | 2.769% | 0.5537% | 0.2132% | 6.115% | 5.207% | -1.262% | -1.376% | 8.839% | -3.759% | 2.324% | -5.802% | -1.217% | -5.309% | 1.235% | -1.489% | -5.121% | 9.973% | -4.099% | 3.277% | 0.3267% | 2.074%
| Z-Score (Absolute) |  | 1.08 | 0.67 | 0.66 | 0.04 | 0.73 | 0.07 | 0.03 | 1.72 | 1.45 | 0.47 | 0.51 | 2.53 | 1.21 | 0.59 | 1.82 | 0.46 | 1.67 | 0.27 | 0.54 | 1.62 | 2.87 | 1.31 | 0.88 | 0 | 0.52
| Analytical results for Zr in OREAS 906b (Certified Value 267 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 269.3 | 276 | 268.3642 | 254.3 | 280.89 | 276 | 257 | 273 | 264 | 266 | 266 | 288 | 256.5 | 272.9112 | 260 | 255.3 | 251.76 | 254 | 267 | 270 | 268.328 | 276 | 274 | 277.9 |  | 1
| 1 | 2 | 265.2 | 278 | 254.5189 | 264.2 | 293.91 | 277 | 253 | 284 | 267 | 257 | 267 | 285 | 255.6 | 273.2225 | 262 | 259.5 | 250.1 | 261 | 263 | 286 | 264.648 | 275 | 172 | 267.4 |  | 14
| 1 | 3 | 270.9 | 275 | 259.57 | 264.3 | 285.52 | 280 | 252 | 288 | 267 | 247 | 267 | 276 | 259.5 | 270.8767 | 255 | 253.9 | 249.48 | 260 | 267 | 270 | 266.864 | 276 | 266 | 266.1 |  | 16
| 1 | 4 | 264 | 284 | 256.0834 | 264.5 | 286.77 | 274 | 250 | 290 | 267 | 254 | 262 | 277 | 262.4 | 268.4724 | 257 | 264.7 | 250.89 | 263 | 271 | 265 | 266.904 | 279 | 290 | 266.5 |  | 267.4545
| 1 | 5 | 261.7 | 279 | 269.8316 | 267.6 | 291.07 | 280 | 256 | 277 | 268 | 261 | 265 | 284 | 258.5 | 268.2224 | 255 | 263.8 | 251.06 | 262 | 277 | 272 | 269.712 | 278 | 277 | 262.2 |  | 77
| 1 | 6 | 269.3 | 271 | 272.2191 | 261.4 | 282.79 | 276 | 252 | 283 | 261 | 258 | 265 | 282 | 258.6 | 275.4638 | 258 | 259.1 | 252.37 | 259 | 264 | 252 | 262.032 | 275 | 263 | 262.3
| Mean |  | 266.7333 | 277.1667 | 263.4312 | 262.7167 | 286.825 | 277.1667 | 253.3333 | 282.5 | 265.6667 | 257.1667 | 265.3333 | 282 | 258.5167 | 271.5281 | 257.8333 | 259.3833 | 250.9433 | 259.8333 | 268.1667 | 269.1667 | 266.4147 | 276.5 | 257 | 267.0667
| Median |  | 267.25 | 277 | 263.9671 | 264.25 | 286.145 | 276.5 | 252.5 | 283.5 | 267 | 257.5 | 265.5 | 283 | 258.55 | 271.8939 | 257.5 | 259.3 | 250.975 | 260.5 | 267 | 270 | 266.884 | 276 | 270 | 266.3
| Std Dev. |  | 3.6247 | 4.3551 | 7.627 | 4.5675 | 4.9281 | 2.4014 | 2.6583 | 6.473 | 2.6583 | 6.4317 | 1.8619 | 4.6904 | 2.3928 | 2.8623 | 2.7869 | 4.3499 | 1.0556 | 3.1885 | 5.1543 | 10.9985 | 2.7314 | 1.6432 | 42.7083 | 5.7469
| Rel.Std.Dev. |  | 1.359% | 1.571% | 2.895% | 1.739% | 1.718% | 0.8664% | 1.049% | 2.291% | 1.001% | 2.501% | 0.7017% | 1.663% | 0.9256% | 1.054% | 1.081% | 1.677% | 0.4206% | 1.227% | 1.922% | 4.086% | 1.025% | 0.5943% | 16.62% | 2.152%
| PDM3 |  | -0.2697% | 3.631% | -1.504% | -1.771% | 7.243% | 3.631% | -5.28% | 5.625% | -0.6685% | -3.847% | -0.7931% | 5.438% | -3.342% | 1.523% | -3.597% | -3.018% | -6.173% | -2.85% | 0.2663% | 0.6402% | -0.3888% | 3.382% | -3.909% | -0.145%
| Z-Score (Absolute) |  | 0.07 | 1.06 | 0.25 | 0.32 | 1.98 | 1.06 | 1.21 | 1.57 | 0.04 | 0.84 | 0.07 | 1.52 | 0.72 | 0.52 | 0.78 | 0.63 | 1.44 | 0.59 | 0.2 | 0.3 | 0.04 | 0.99 | 0.86 | 0.1
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C222: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C240: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C259: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C277: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C295: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 906b (Indicative Value 0.621 ppm) | [21 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.06 | 2.12 | 2.20 | [17 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES | [17 empty] | 3
3
| 1 | 1 | 0.56 | < 1 | 0.7 | [17 empty] | 1
| 1 | 2 | 0.57 | < 1 | 0.7 | [17 empty] | 9
| 1 | 3 | 0.59 | < 1 | 0.5 | [17 empty] | 16
| 1 | 4 | 0.58 | < 1 | 0.5 | [17 empty] | 0.62083
| 1 | 5 | 0.6 | < 1 | 0.8 | [17 empty] | 15
| 1 | 6 | 0.55 | < 1 | 0.8
| Mean |  | 0.575 |  | 0.66667
| Median |  | 0.575 |  | 0.7
| Std Dev. |  | 0.018708 |  | 0.13663
| Rel.Std.Dev. |  | 3.254% |  | 20.49%
| PDM3 |  | -7.383% |  | 7.383%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.82
| Analytical results for Al in OREAS 906b (Certified Value 7.57 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 2
3
| 1 | 1 | 7.54 | 7.76 | 7.3324 | 7.67 | 3.59 | 7.621 | 7.383 | 7.33 | 6.854 | 7.7 | 7.77 | 8.1 | 7.4 | 7.05 | 7.73 | 7.551 | 7.66 | 7.41 |  |  | 1
| 1 | 2 | 7.61 | 7.55 | 7.2454 | 7.59 | 3.57 | 7.542 | 7.383 | 7.277 | 7.198 | 7.727 | 7.81 | 8.17 | 7.51 | 7.35 | 7.7 | 7.586 | 7.5 | 7.41 |  |  | #N/A
| 1 | 3 | 7.73 | 7.45 | 7.2092 | 7.95 | 3.6 | 7.489 | 7.304 | 7.277 | 7.145 | 7.859 | 7.55 | 7.89 | 7.41 | 7.01 | 7.7 | 7.566 | 7.62 | 7.5 |  |  | 16
| 1 | 4 | 7.8 | 7.64 | 7.1529 | 7.75 | 3.57 | 7.621 | 7.489 | 7.171 | 6.377 | 7.489 | 8.15 | 8.07 | 7.56 | 7.35 | 7.71 | 7.508 | 7.65 | 7.52 |  |  | 7.571
| 1 | 5 | 7.67 | 7.53 | 7.3567 | 7.81 | 3.58 | 7.595 | 7.383 | 7.436 | 6.616 | 7.674 | 7.86 | 7.98 | 7.41 | 7.38 | 7.72 | 7.561 | 7.64 | 7.52 |  |  | 79
| 1 | 6 | 7.53 | 7.62 | 7.0892 | 7.83 | 3.58 | 7.621 | 7.409 | 7.462 | 6.642 | 7.648 | 8 | 7.81 | 7.54 | 7.13 | 7.73 | 7.592 | 7.61 | 7.56
| Mean |  | 7.6467 | 7.5917 | 7.231 | 7.7667 | 3.5817 | 7.5815 | 7.3918 | 7.3255 | 6.8053 | 7.6828 | 7.8567 | 8.0033 | 7.4717 | 7.2117 | 7.715 | 7.5607 | 7.6133 | 7.4867
| Median |  | 7.64 | 7.585 | 7.2273 | 7.78 | 3.58 | 7.608 | 7.383 | 7.3035 | 6.748 | 7.687 | 7.835 | 8.025 | 7.46 | 7.24 | 7.715 | 7.5635 | 7.63 | 7.51
| Std Dev. |  | 0.10708 | 0.10685 | 0.10291 | 0.12675 | 0.01169 | 0.054735 | 0.059472 | 0.10903 | 0.32185 | 0.12012 | 0.2051 | 0.13589 | 0.073052 | 0.16738 | 0.013784 | 0.030064 | 0.058538 | 0.062503
| Rel.Std.Dev. |  | 1.4% | 1.407% | 1.423% | 1.632% | 0.3264% | 0.722% | 0.8046% | 1.488% | 4.729% | 1.563% | 2.611% | 1.698% | 0.9777% | 2.321% | 0.1787% | 0.3976% | 0.7689% | 0.8349%
| PDM3 |  | 0.9991% | 0.2726% | -4.492% | 2.584% | -52.69% | 0.1383% | -2.367% | -3.243% | -10.11% | 1.477% | 3.773% | 5.71% | -1.312% | -4.747% | 1.902% | -0.1368% | 0.5588% | -1.114%
| Z-Score (Absolute) |  | 0.32 | 0.09 | 1.42 | 0.82 | 16.64 | 0.04 | 0.75 | 1.02 | 3.2 | 0.47 | 1.19 | 1.8 | 0.41 | 1.5 | 0.6 | 0.04 | 0.18 | 0.35
| Analytical results for As in OREAS 906b (Indicative Value 14.4 ppm) | [21 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 1
1
| 1 | 1 | < 20 | < 20 | 17 | 9 | < 100 | 100 | 100 | 100 | < 100 | 13 | < 30 | 13 | 22 | < 30 | 15 | < 100 |  |  |  |  | 1
| 1 | 2 | < 20 | < 20 | 17 | 9.3 | < 100 | 100 | 100 | 100 | < 100 | 14 | < 30 | 13 | 21 | < 30 | 14 | < 100 |  |  |  |  | 10
| 1 | 3 | < 20 | < 20 | 15 | 8.5 | < 100 | < 100 | 100 | 100 | < 100 | 13 | < 30 | 12 | 22 | < 30 | 16 | < 100 |  |  |  |  | 16
| 1 | 4 | < 20 | < 20 | < 10 | 8.2 | < 100 | 100 | 100 | < 100 | < 100 | 13 | < 30 | 13 | 21 | < 30 | 16 | < 100 |  |  |  |  | 14.4028
| 1 | 5 | < 20 | < 20 | 15 | 6.9 | 100 | 100 | < 100 | < 100 | < 100 | 13 | < 30 | 13 | 22 | < 30 | 14 | < 100 |  |  |  |  | 16
| 1 | 6 | < 20 | < 20 | 13 | 9.2 | 100 | 100 | < 100 | 100 | < 100 | 12 | < 30 | 12 | 22 | < 30 | 16 | < 100
| Mean |  |  |  | 15.4 | 8.5167 | 100 | 100 | 100 | 100 |  | 13 |  | 12.6667 | 21.6667 |  | 15.1667
| Median |  |  |  | 15 | 8.75 | 100 | 100 | 100 | 100 |  | 13 |  | 13 | 22 |  | 15.5
| Std Dev. |  |  |  | 1.6733 | 0.89759 | 0 | 0 | 0 | 0 |  | 0.63246 |  | 0.5164 | 0.5164 |  | 0.98319
| Rel.Std.Dev. |  |  |  | 10.87% | 10.54% | 0% | 0% | 0% | 0% |  | 4.865% |  | 4.077% | 2.383% |  | 6.483%
| PDM3 |  |  |  | 6.924% | -40.87% | 594.3% | 594.3% | 594.3% | 594.3% |  | -9.74% |  | -12.05% | 50.43% |  | 5.304%
| Z-Score (Absolute) |  | 0.59 | 0.59 | 0.16 | 0.76 | 5.97 | 8.87 | 7.9 | 7.9 | 4.04 | 0.24 | 0.01 | 0.28 | 0.76 | 0.01 | 0.01 | 4.04
| Analytical results for B in OREAS 906b (Indicative Value 41.4 ppm) | [21 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 2.03 | 2.20 | 2.25 | 2.26 | [16 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*MS | [16 empty] | 1
1
| 1 | 1 | < 50 | 35 | < 10 | 54 | [16 empty] | 1
| 1 | 2 | < 50 | 30 | < 10 | 49 | [16 empty] | 7
| 1 | 3 | < 50 | 37 | < 10 | 55 | [16 empty] | 16
| 1 | 4 | < 50 | 35 | < 10 | 49 | [16 empty] | 41.4167
| 1 | 5 | < 50 | 29 | < 10 | 62 | [16 empty] | 17
| 1 | 6 | < 50 | 23 | < 10 | 39
| Mean |  |  | 31.5 |  | 51.3333
| Median |  |  | 32.5 |  | 51.5
| Std Dev. |  |  | 5.2058 |  | 7.7115
| Rel.Std.Dev. |  |  | 16.53% |  | 15.02%
| PDM3 |  |  | -23.94% |  | 23.94%
| Z-Score (Absolute) |  | 0.17 | 0.17 | 1.19 | 1.18
| Analytical results for Ba in OREAS 906b (Certified Value 2841 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | 2846 | 3160 | 2856.0744 | 2900 | 195.2 | 2913 | 2814.2 | 3103 | 2610 | 2833 | 2428 | 2830 | [8 empty] | 1
| 1 | 2 | 2840 | 3140 | 2742.7575 | 2900 | 193.2 | 2909 | 2809 | 3125 | 2710 | 2831 | 2476 | 2770 | [8 empty] | 35
| 1 | 3 | 2822 | 3220 | 2691.1507 | 2900 | 196.4 | 2851 | 2803.8 | 3085 | 2620 | 2832 | 2463 | 2870 | [8 empty] | 16
| 1 | 4 | 2816 | 3240 | 2759.8571 | 2900 | 193.9 | 2779 | 2848.7 | 3130 | 2670 | 2834 | 2458 | 2850 | [8 empty] | 2841.4353
| 1 | 5 | 2855 | 3260 | 2847.4245 | 2900 | 194.7 | 2801 | 2820.2 | 3043 | 2670 | 2832 | 2421 | 2860 | [8 empty] | 80
| 1 | 6 | 2830 | 3210 | 2691.8405 | 2900 | 196.3 | 2772 | 2841.5 | 3063 | 2600 | 2836 | 2433 | 2870
| Mean |  | 2834.8333 | 3205 | 2764.8508 | 2900 | 194.95 | 2837.5 | 2822.9 | 3091.5 | 2646.6667 | 2833 | 2446.5 | 2841.6667
| Median |  | 2835 | 3215 | 2751.3073 | 2900 | 194.95 | 2826 | 2817.2 | 3094 | 2645 | 2832.5 | 2445.5 | 2855
| Std Dev. |  | 14.838 | 46.3681 | 72.6848 | 0 | 1.2818 | 63.308 | 18.1802 | 34.4775 | 43.2049 | 1.7889 | 22.1337 | 38.1663
| Rel.Std.Dev. |  | 0.5234% | 1.447% | 2.629% | 0% | 0.6575% | 2.231% | 0.644% | 1.115% | 1.632% | 0.06314% | 0.9047% | 1.343%
| PDM3 |  | -0.2323% | 12.8% | -2.695% | 2.061% | -93.14% | -0.1385% | -0.6523% | 8.801% | -6.855% | -0.2969% | -13.9% | 0.008144%
| Z-Score (Absolute) |  | 0.01 | 3.7 | 0.69 | 0.66 | 26.33 | 0.04 | 0.11 | 2.57 | 1.87 | 0.01 | 3.87 | 0.08
| Analytical results for Be in OREAS 906b (Certified Value 3.15 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 3 | 3 | 3.1869 | 4 | < 2 | < 5 | < 5 | 2.92 | 3 | 7 | < 5 | < 3 | NA | [7 empty] | 1
| 1 | 2 | 3 | 3 | 3.3238 | 4 | < 2 | < 5 | < 5 | 3.01 | 3 | 6 | < 5 | < 3 | NA | [7 empty] | 36
| 1 | 3 | 3 | 3 | 3.0473 | < 2 | < 2 | < 5 | < 5 | 2.87 | 3 | 7 | < 5 | 4 | NA | [7 empty] | 16
| 1 | 4 | 4 | 3 | 3.2565 | 2 | < 2 | < 5 | < 5 | 3.06 | 3 | 6 | < 5 | < 3 | NA | [7 empty] | 3.1472
| 1 | 5 | 3 | 3 | 3.2125 | 3 | < 2 | < 5 | < 5 | 3.07 | 3 | 7 | < 5 | 3 | NA | [7 empty] | 81
| 1 | 6 | 4 | 3 | 2.9949 | 3 | < 2 | < 5 | < 5 | 3.03 | 3 | 7 | < 5 | 3 | NA
| Mean |  | 3.3333 | 3 | 3.1703 | 3.2 |  |  |  | 2.9933 | 3 | 6.6667 |  | 3.3333
| Median |  | 3 | 3 | 3.1997 | 3 |  |  |  | 3.02 | 3 | 7 |  | 3
| Std Dev. |  | 0.5164 | 0 | 0.12567 | 0.83666 |  |  |  | 0.080664 | 0 | 0.5164 |  | 0.57735
| Rel.Std.Dev. |  | 15.49% | 0% | 3.964% | 26.15% |  |  |  | 2.695% | 0% | 7.746% |  | 17.32%
| PDM3 |  | 5.915% | -4.677% | 0.7344% | 1.678% |  |  |  | -4.889% | -4.677% | 111.8% |  | 5.915%
| Z-Score (Absolute) |  | 0.69 | 0.14 | 0.42 | 0.13 | 3.12 | 0.67 | 0.67 | 0.13 | 0.14 | 6.12 | 0.67 | 0.81 | NA
| Analytical results for Bi in OREAS 906b (Certified Value 10.4 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.20 | 2.25 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES/MS | [12 empty] | 1
2
| 1 | 1 | 10.3 | 10.4 | 10.6167 | < 5 | 11.1 | 9.6 | 10.2 | 10 | [12 empty] | 1
| 1 | 2 | 10.5 | 10.4 | 10.2487 | < 5 | 11 | 9.3 | 10.8 | 10 | [12 empty] | 37
| 1 | 3 | 10.4 | 10.5 | 9.9241 | < 5 | 11 | 9.8 | 11.1 | 10 | [12 empty] | 16
| 1 | 4 | 10.2 | 10.5 | 10.2302 | < 5 | 10.6 | 9.6 | 10.3 | 10 | [12 empty] | 10.4407
| 1 | 5 | 10.2 | 10.5 | 10.363 | < 5 | 11 | 9 | 10.3 | 10 | [12 empty] | 82
| 1 | 6 | 10.2 | 10.5 | 9.7375 | < 5 | 10.7 | 9.6 | 9.4 | 10
| Mean |  | 10.3 | 10.4667 | 10.1867 |  | 10.9 | 9.4833 | 10.35 | 10
| Median |  | 10.25 | 10.5 | 10.2395 |  | 11 | 9.6 | 10.3 | 10
| Std Dev. |  | 0.12649 | 0.05164 | 0.31387 |  | 0.2 | 0.28577 | 0.58224 | 0
| Rel.Std.Dev. |  | 1.228% | 0.4934% | 3.081% |  | 1.835% | 3.013% | 5.625% | 0%
| PDM3 |  | -1.347% | 0.249% | -2.433% |  | 4.399% | -9.169% | -0.8685% | -4.221%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.46 | 31.56 | 2.43 | 3.3 | 0.2 | NA
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 906b (Certified Value 0.638 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.6 | 0.8 | 0.6516 | 0.72 | 0.6 | 0.622 | 0.693 | 0.7 | 0.643 | 0.65 | 0.7 | 1.19 | 0.65701 | 0.658 | 0.56 | 0.58 | 0.65 | 0.68 | 0.61 |  | 1
| 1 | 2 | 0.6 | 0.8 | 0.63111 | 0.67 | 0.6 | 0.643 | 0.672 | 0.65 | 0.7 | 0.686 | 0.7 | 1.18 | 0.65156 | 0.65 | 0.6 | 0.59 | 0.683 | 0.56 | 0.64 |  | #N/A
| 1 | 3 | 0.7 | 0.8 | 0.58916 | 0.78 | 0.59 | 0.586 | 0.672 | 0.643 | 0.729 | 0.679 | 0.6 | 1.2 | 0.65864 | 0.636 | 0.57 | 0.59 | 0.667 | 0.66 | 0.61 |  | 16
| 1 | 4 | 0.7 | 0.8 | 0.62377 | 0.72 | 0.6 | 0.6 | 0.672 | 0.658 | 0.572 | 0.686 | 0.6 | 1.18 | 0.6529 | 0.658 | 0.59 | 0.58 | 0.684 | 0.62 | 0.62 |  | 0.63822
| 1 | 5 | 0.6 | 0.8 | 0.62666 | 0.78 | 0.59 | 0.629 | 0.679 | 0.65 | 0.607 | 0.65 | 0.6 | 1.1 | 0.65792 | 0.665 | 0.59 | 0.58 | 0.662 | 0.66 | 0.61 |  | 83
| 1 | 6 | 0.6 | 0.8 | 0.59016 | 0.7 | 0.59 | 0.586 | 0.693 | 0.758 | 0.6 | 0.643 | 0.6 | 1.14 | 0.65068 | 0.672 | 0.66 | 0.58 | 0.661 | 0.62 | 0.63
| Mean |  | 0.63333 | 0.8 | 0.61875 | 0.72833 | 0.595 | 0.611 | 0.68017 | 0.6765 | 0.64183 | 0.66567 | 0.63333 | 1.165 | 0.65478 | 0.6565 | 0.595 | 0.58333 | 0.66783 | 0.63333 | 0.62
| Median |  | 0.6 | 0.8 | 0.62522 | 0.72 | 0.595 | 0.611 | 0.6755 | 0.654 | 0.625 | 0.6645 | 0.6 | 1.18 | 0.65495 | 0.658 | 0.59 | 0.58 | 0.6645 | 0.64 | 0.615
| Std Dev. |  | 0.05164 | 1.2162e-16 | 0.024545 | 0.044008 | 0.0054772 | 0.023833 | 0.010304 | 0.044863 | 0.061356 | 0.020047 | 0.05164 | 0.037815 | 0.003477 | 0.012486 | 0.035071 | 0.005164 | 0.013348 | 0.043205 | 0.012649
| Rel.Std.Dev. |  | 8.154% | 1.52e-14% | 3.967% | 6.042% | 0.9205% | 3.901% | 1.515% | 6.632% | 9.56% | 3.012% | 8.154% | 3.246% | 0.531% | 1.902% | 5.894% | 0.8853% | 1.999% | 6.822% | 2.04%
| PDM3 |  | -0.7661% | 25.35% | -3.052% | 14.12% | -6.772% | -4.265% | 6.572% | 5.997% | 0.5657% | 4.3% | -0.7661% | 82.54% | 2.595% | 2.864% | -6.772% | -8.6% | 4.64% | -0.7661% | -2.855%
| Z-Score (Absolute) |  | NA | NA | 0.69 | 1.86 | 1.24 | 0.87 | 0.74 | 0.65 | 0.15 | 0.41 | NA | 12 | 0.15 | 0.18 | 1.24 | 1.51 | 0.45 | 0.35 | 0.66
Results from laboratories 2.01, 2.02 and 2.12 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 906b (Indicative Value 0.33 ppm) | [21 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.20 | 2.24 | 2.25 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 10 | < 5 | < 1 | 0.4 | 0.3 | 0.2 | < 10 | < 2 | [12 empty] | 1
| 1 | 2 | < 10 | < 5 | < 1 | 0.3 | 0.4 | 0.2 | < 10 | < 2 | [12 empty] | 12
| 1 | 3 | < 10 | < 5 | < 1 | 0.4 | 0.4 | 0.2 | < 10 | < 2 | [12 empty] | 16
| 1 | 4 | < 10 | < 5 | < 1 | 0.5 | 0.3 | 0.3 | < 10 | < 2 | [12 empty] | 0.33333
| 1 | 5 | < 10 | 5 | < 1 | 0.4 | 0.3 | 0.2 | < 10 | < 2 | [12 empty] | 18
| 1 | 6 | < 10 | < 5 | < 1 | 0.4 | 0.4 | 0.4 | < 10 | < 2
| Mean |  |  | 5 |  | 0.4 | 0.35 | 0.25
| Median |  |  | 5 |  | 0.4 | 0.35 | 0.2
| Std Dev. |  |  |  |  | 0.063246 | 0.054772 | 0.083666
| Rel.Std.Dev. |  |  |  |  | 15.81% | 15.65% | 33.47%
| PDM3 |  |  | 1400% |  | 20% | 5% | -25%
| Z-Score (Absolute) |  | 6.37 | 3.25 | 0.37 | 0.52 | 0.6 | 0.75 | 6.37 | 0.37
| Analytical results for Ce in OREAS 906b (Certified Value 88 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [11 empty] | 0
1
| 1 | 1 | 90.6 | 81 | 92.5 | 93.3 | 61.6 | 77.36 | 95.8 | 90.9 | 95 | [11 empty] | 1
| 1 | 2 | 88.5 | 82 | 84 | 96.7 | 62.8 | 79.11 | 95.2 | 83.1 | 98 | [11 empty] | 39
| 1 | 3 | 89.3 | 77 | 83.2 | 93 | 62.4 | 78.29 | 86.6 | 88.4 | 93 | [11 empty] | 16
| 1 | 4 | 88.5 | 84 | 86.8 | 91.5 | 64.1 | 78.33 | 84.6 | 88.8 | 93 | [11 empty] | 87.7798
| 1 | 5 | 89.9 | 82 | 90 | 95.1 | 63.4 | 77.26 | 94.5 | 88.9 | 94 | [11 empty] | 84
| 1 | 6 | 88.2 | 84 | 85.7 | 93.9 | 61.5 | 78.18 | 82.1 | 92.5 | 94
| Mean |  | 89.1667 | 81.6667 | 87.0333 | 93.9167 | 62.6333 | 78.0883 | 89.8 | 88.7667 | 94.5
| Median |  | 88.9 | 82 | 86.25 | 93.6 | 62.6 | 78.235 | 90.55 | 88.85 | 94
| Std Dev. |  | 0.94163 | 2.582 | 3.5892 | 1.8005 | 1.0172 | 0.68817 | 6.0633 | 3.186 | 1.8708
| Rel.Std.Dev. |  | 1.056% | 3.162% | 4.124% | 1.917% | 1.624% | 0.8813% | 6.752% | 3.589% | 1.98%
| PDM3 |  | 1.58% | -6.964% | -0.8504% | 6.991% | -28.65% | -11.04% | 2.301% | 1.124% | 7.656%
| Z-Score (Absolute) |  | 0.05 | 0.93 | 0.23 | 0.67 | 3.42 | 1.4 | 0.14 | 0 | 0.75
| Analytical results for Co in OREAS 906b (Certified Value 17.5 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  | 1
1
| 1 | 1 | 18 | 18.423 | 16 | 14.8 | < 20 | < 20 | < 20 | < 20 | < 20 | 18.1 | 17.6 | 19 | 16 | 20 | 14 | 20.3 | < 20 |  |  |  | 1
| 1 | 2 | 18 | 17.383 | 16 | 13.7 | < 20 | < 20 | 20 | < 20 | < 20 | 18.6 | 17.7 | 20.1 | 17 | 19 | 14 | 18 | < 20 |  |  |  | 27
| 1 | 3 | 18 | 16.6559 | 14 | 15.2 | < 20 | 20 | < 20 | < 20 | < 20 | 18 | 17.4 | 19.2 | 17 | 19 | 14 | 18.9 | < 20 |  |  |  | 16
| 1 | 4 | 17 | 18.144 | 17 | 14.4 | < 20 | 30 | < 20 | < 20 | < 20 | 17.8 | 17.5 | 19.2 | 17 | 19 | 15 | 17.7 | < 20 |  |  |  | 17.5452
| 1 | 5 | 18 | 17.5394 | 15 | 14.1 | < 20 | < 20 | 20 | < 20 | < 20 | 18 | 17.3 | 19.4 | 16 | 19 | 13 | 17.9 | < 20 |  |  |  | 85
| 1 | 6 | 17 | 16.6765 | 16 | 13.3 | 30 | 20 | < 20 | < 20 | < 20 | 18 | 17.2 | 19.6 | 18 | 20 | 14 | 17.6 | < 20
| Mean |  | 17.6667 | 17.4703 | 15.6667 | 14.25 | 30 | 23.3333 | 20 |  |  | 18.0833 | 17.45 | 19.4167 | 16.8333 | 19.3333 | 14 | 18.4
| Median |  | 18 | 17.4612 | 16 | 14.25 | 30 | 20 | 20 |  |  | 18 | 17.45 | 19.3 | 17 | 19 | 14 | 17.95
| Std Dev. |  | 0.5164 | 0.73036 | 1.0328 | 0.70071 |  | 5.7735 | 0 |  |  | 0.27142 | 0.18708 | 0.392 | 0.75277 | 0.5164 | 0.63246 | 1.0392
| Rel.Std.Dev. |  | 2.923% | 4.181% | 6.592% | 4.917% |  | 24.74% | 0% |  |  | 1.501% | 1.072% | 2.019% | 4.472% | 2.671% | 4.518% | 5.648%
| PDM3 |  | 0.6926% | -0.4266% | -10.71% | -18.78% | 70.99% | 32.99% | 13.99% |  |  | 3.067% | -0.5423% | 10.67% | -4.057% | 10.19% | -20.21% | 4.872%
| Z-Score (Absolute) |  | 0.28 | 0.22 | 0.28 | 0.67 | 0.93 | 0 | 0.93 | 1.86 | 1.86 | 0.4 | 0.22 | 0.77 | 0.05 | 0.74 | 0.74 | 0.48 | 1.86
| Analytical results for Cr in OREAS 906b (Indicative Value 34.3 ppm) | [21 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 1
1
| 1 | 1 | < 50 | < 50 | < 100 | 70 | < 70 | 140 | 70 | < 70 | 11 | 47 | < 70 | 10 | 16 | 10 | 40 | < 100 |  |  |  |  | 1
| 1 | 2 | < 50 | < 50 | < 100 | 70 | < 70 | 70 | 70 | < 70 | 15 | 46 | < 70 | 13 | 16 | 11 | < 30 | < 100 |  |  |  |  | 13
| 1 | 3 | < 50 | < 50 | < 100 | < 70 | 70 | 70 | < 70 | < 70 | 10 | 48 | < 70 | 13 | 16 | 11 | < 30 | < 100 |  |  |  |  | 16
| 1 | 4 | < 50 | < 50 | < 100 | < 70 | 70 | 70 | 70 | < 70 | < 10 | 45 | < 70 | 15 | 16 | 12 | < 30 | < 100 |  |  |  |  | 34.3304
| 1 | 5 | < 50 | < 50 | < 100 | < 70 | 70 | 70 | 70 | < 70 | < 10 | 44 | < 70 | 12 | 16 | 11 | < 30 | < 100 |  |  |  |  | 19
| 1 | 6 | < 50 | < 50 | < 100 | < 70 | 70 | 140 | 140 | < 70 | 20 | 41 | < 70 | 13 | 16 | 10 | < 30 | < 100
| Mean |  |  |  |  | 70 | 70 | 93.3333 | 84 |  | 14 | 45.1667 |  | 12.6667 | 16 | 10.8333 | 40
| Median |  |  |  |  | 70 | 70 | 70 | 70 |  | 13 | 45.5 |  | 13 | 16 | 11 | 40
| Std Dev. |  |  |  |  | 0 | 0 | 36.1478 | 31.305 |  | 4.5461 | 2.4833 |  | 1.633 | 0 | 0.75277
| Rel.Std.Dev. |  |  |  |  | 0% | 0% | 38.73% | 37.27% |  | 32.47% | 5.498% |  | 12.89% | 0% | 6.949%
| PDM3 |  |  |  |  | 103.9% | 103.9% | 171.9% | 144.7% |  | -59.22% | 31.56% |  | -63.1% | -53.39% | -68.44% | 16.51%
| Z-Score (Absolute) |  | 0.39 | 0.39 | 0.67 | 0.49 | 0.97 | 2.44 | 1.7 | 0 | 0.99 | 0.47 | 0 | 0.92 | 0.78 | 1 | 0.64 | 0.67
| Analytical results for Cs in OREAS 906b (Certified Value 7.05 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 7.6 | 7.25 | 6.8991 | 6.9 | 6.8 | 7.54 | 6.6 | 7.4 | [12 empty] | 1
| 1 | 2 | 7.2 | 7.31 | 6.7832 | 7.1 | 6.7 | 7.59 | 6.7 | 7 | [12 empty] | 42
| 1 | 3 | 7.1 | 7.27 | 6.6404 | 6.8 | 6.2 | 7.62 | 6.3 | 7.2 | [12 empty] | 16
| 1 | 4 | 7.6 | 7.2 | 7.0471 | 6.8 | 6.4 | 7.52 | 6.6 | 7.8 | [12 empty] | 7.049
| 1 | 5 | 7.2 | 7.21 | 6.6091 | 7.2 | 6.7 | 7.61 | 6.4 | 7.9 | [12 empty] | 86
| 1 | 6 | 7.1 | 7.4 | 6.4232 | 7 | 6.6 | 7.63 | 6.9 | 7
| Mean |  | 7.3 | 7.2733 | 6.7337 | 6.9667 | 6.5667 | 7.585 | 6.5833 | 7.3833
| Median |  | 7.2 | 7.26 | 6.7118 | 6.95 | 6.65 | 7.6 | 6.6 | 7.3
| Std Dev. |  | 0.23664 | 0.073937 | 0.22301 | 0.1633 | 0.22509 | 0.045056 | 0.2137 | 0.392
| Rel.Std.Dev. |  | 3.242% | 1.017% | 3.312% | 2.344% | 3.428% | 0.594% | 3.246% | 5.309%
| PDM3 |  | 3.561% | 3.182% | -4.473% | -1.168% | -6.843% | 7.604% | -6.606% | 4.743%
| Z-Score (Absolute) |  | 0.37 | 0.32 | 0.8 | 0.32 | 1.15 | 0.97 | 1.11 | 0.55
| Analytical results for Cu in OREAS 906b (Certified Value 0.311 wt.%) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | 3
3
| 1 | 1 | 0.3235 | 0.307 | 0.31158 | 0.3301 | 0.32641 | 0.313 | 0.314 | 0.31 | 0.287 | 0.298 | 0.3215 | 0.3166 | 0.28985 | 0.22823 | 0.278 | 0.337 | 0.3058 | 0.319 | 0.3236 | 0.304 | 1
| 1 | 2 | 0.3201 | 0.302 | 0.31227 | 0.324 | 0.32101 | 0.304 | 0.313 | 0.325 | 0.302 | 0.305 | 0.3244 | 0.3221 | 0.28284 | 0.22537 | 0.279 | 0.3388 | 0.3056 | 0.315 | 0.3241 | 0.314 | 29
| 1 | 3 | 0.3199 | 0.306 | 0.30653 | 0.3338 | 0.32676 | 0.307 | 0.308 | 0.305 | 0.314 | 0.305 | 0.3141 | 0.3219 | 0.2815 | 0.23844 | 0.27 | 0.339 | 0.3036 | 0.313 | 0.3188 | 0.305 | 16
| 1 | 4 | 0.3228 | 0.307 | 0.30314 | 0.3298 | 0.33243 | 0.312 | 0.315 | 0.303 | 0.267 | 0.292 | 0.3402 | 0.32 | 0.28368 | 0.24522 | 0.274 | 0.3365 | 0.3062 | 0.323 | 0.3148 | 0.308 | 0.31092
| 1 | 5 | 0.3197 | 0.306 | 0.30148 | 0.3296 | 0.33222 | 0.31 | 0.313 | 0.309 | 0.279 | 0.297 | 0.3236 | 0.3124 | 0.28862 | 0.22309 | 0.27 | 0.336 | 0.3042 | 0.314 | 0.3361 | 0.308 | 87
| 1 | 6 | 0.3231 | 0.309 | 0.31128 | 0.3326 | 0.32234 | 0.305 | 0.312 | 0.313 | 0.28 | 0.297 | 0.3353 | 0.316 | 0.2883 | 0.24105 | 0.272 | 0.3364 | 0.3035 | 0.319 | 0.3328 | 0.307
| Mean |  | 0.32152 | 0.30617 | 0.30771 | 0.32998 | 0.32686 | 0.3085 | 0.3125 | 0.31083 | 0.28817 | 0.299 | 0.32652 | 0.31817 | 0.2858 | 0.23357 | 0.27383 | 0.33728 | 0.30482 | 0.31717 | 0.32503 | 0.30767
| Median |  | 0.32145 | 0.3065 | 0.3089 | 0.32995 | 0.32659 | 0.3085 | 0.313 | 0.3095 | 0.2835 | 0.2975 | 0.324 | 0.3183 | 0.28599 | 0.23334 | 0.273 | 0.33675 | 0.3049 | 0.317 | 0.32385 | 0.3075
| Std Dev. |  | 0.0017893 | 0.0023166 | 0.0046799 | 0.0033885 | 0.0047915 | 0.0037283 | 0.002429 | 0.0078081 | 0.017081 | 0.005099 | 0.0095573 | 0.0038255 | 0.0035308 | 0.0091758 | 0.00392 | 0.0012937 | 0.0011907 | 0.0038166 | 0.0081153 | 0.0035024
| Rel.Std.Dev. |  | 0.5565% | 0.7566% | 1.521% | 1.027% | 1.466% | 1.209% | 0.7773% | 2.512% | 5.928% | 1.705% | 2.927% | 1.202% | 1.235% | 3.929% | 1.432% | 0.3836% | 0.3906% | 1.203% | 2.497% | 1.138%
| PDM3 |  | 3.407% | -1.529% | -1.033% | 6.131% | 5.127% | -0.779% | 0.5075% | -0.02855% | -7.319% | -3.834% | 5.016% | 2.33% | -8.081% | -24.88% | -11.93% | 8.478% | -1.964% | 2.008% | 4.539% | -1.047%
| Z-Score (Absolute) |  | 0.71 | 0.21 | 0.12 | 1.22 | 1.03 | 0.07 | 0.17 | 0.07 | 1.29 | 0.64 | 1.01 | 0.51 | 1.43 | 4.56 | 2.15 | 1.65 | 0.29 | 0.45 | 0.92 | 0.12
| Analytical results for Dy in OREAS 906b (Certified Value 3.74 ppm) | [21 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 3.6 | 3.7 | 3.81 | 3.99 | 3.29 | 3.86 | 3.5 | 4 | [12 empty] | 1
| 1 | 2 | 3.7 | 3.9 | 3.91 | 4.12 | 3.55 | 4.09 | 3.1 | 4 | [12 empty] | 16
| 1 | 3 | 3.8 | 3.5 | 3.87 | 3.96 | 3.39 | 3.81 | 3.9 | 4 | [12 empty] | 16
| 1 | 4 | 3.7 | 3.8 | 3.79 | 4.02 | 3.28 | 3.74 | 3.5 | 4 | [12 empty] | 3.7447
| 1 | 5 | 3.7 | 4 | 3.79 | 3.88 | 3.28 | 4.02 | 3.9 | 4 | [12 empty] | 88
| 1 | 6 | 3.7 | 3.8 | 3.8 | 3.94 | 3.3 | 3.47 | 4.5 | 4
| Mean |  | 3.7 | 3.7833 | 3.8283 | 3.985 | 3.3483 | 3.8317 | 3.7333 | 4
| Median |  | 3.7 | 3.8 | 3.805 | 3.975 | 3.295 | 3.835 | 3.7 | 4
| Std Dev. |  | 0.063246 | 0.17224 | 0.049967 | 0.081425 | 0.10722 | 0.22031 | 0.48028 | 0
| Rel.Std.Dev. |  | 1.709% | 4.553% | 1.305% | 2.043% | 3.202% | 5.75% | 12.86% | 0%
| PDM3 |  | -1.193% | 1.033% | 2.234% | 6.418% | -10.58% | 2.323% | -0.3027% | 6.819%
| Z-Score (Absolute) |  | 1.12 | 0 | 0.61 | 2.72 | 5.87 | 0.65 | 0.67 | NA
| Analytical results for Er in OREAS 906b (Certified Value 1.13 ppm) | [21 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 1.2 | 1.1 | 1.14 | 1.07 | 0.92 | 1.09 | 1.3 | < 2 | [12 empty] | 1
| 1 | 2 | 1.1 | 1.2 | 1.24 | 1.07 | 0.97 | 1.18 | 0.9 | < 2 | [12 empty] | 17
| 1 | 3 | 1.1 | 1 | 1.28 | 1.11 | 0.94 | 1.08 | 1.2 | < 2 | [12 empty] | 16
| 1 | 4 | 1.1 | 1.1 | 1.2 | 1.17 | 0.93 | 1.07 | 1 | < 2 | [12 empty] | 1.1261
| 1 | 5 | 1 | 1.3 | 1.2 | 1.06 | 0.95 | 1.14 | 1.2 | < 2 | [12 empty] | 89
| 1 | 6 | 1.1 | 1.1 | 1.2 | 1.08 | 0.95 | 0.95 | 1.3 | < 2
| Mean |  | 1.1 | 1.1333 | 1.21 | 1.0933 | 0.94333 | 1.085 | 1.15
| Median |  | 1.1 | 1.1 | 1.2 | 1.075 | 0.945 | 1.085 | 1.2
| Std Dev. |  | 0.063246 | 0.10328 | 0.046904 | 0.041312 | 0.017512 | 0.078166 | 0.16432
| Rel.Std.Dev. |  | 5.75% | 9.113% | 3.876% | 3.779% | 1.856% | 7.204% | 14.29%
| PDM3 |  | -2.314% | 0.6463% | 7.455% | -2.906% | -16.23% | -3.646% | 2.126%
| Z-Score (Absolute) |  | 0.05 | 0.55 | 1.7 | 0.05 | 2.3 | 0.17 | 0.8 | 1.45
| Analytical results for Eu in OREAS 906b (Certified Value 1.51 ppm) | [21 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 1.5 | 1.5 | 1.48 | 1.54 | 1.24 | 1.61 | 1.6 | < 2 | [12 empty] | 1
| 1 | 2 | 1.4 | 1.4 | 1.56 | 1.48 | 1.28 | 1.64 | 1.4 | < 2 | [12 empty] | 18
| 1 | 3 | 1.5 | 1.5 | 1.56 | 1.45 | 1.23 | 1.58 | 1.5 | < 2 | [12 empty] | 16
| 1 | 4 | 1.4 | 1.5 | 1.46 | 1.46 | 1.28 | 1.62 | 1.6 | < 2 | [12 empty] | 1.5121
| 1 | 5 | 1.4 | 1.5 | 1.56 | 1.46 | 1.25 | 1.93 | 1.7 | < 2 | [12 empty] | 90
| 1 | 6 | 1.4 | 1.5 | 1.58 | 1.47 | 1.23 | 1.51 | 1.6 | < 2
| Mean |  | 1.4333 | 1.4833 | 1.5333 | 1.4767 | 1.2517 | 1.6483 | 1.5667
| Median |  | 1.4 | 1.5 | 1.56 | 1.465 | 1.245 | 1.615 | 1.6
| Std Dev. |  | 0.05164 | 0.040825 | 0.050067 | 0.03266 | 0.023166 | 0.14525 | 0.10328
| Rel.Std.Dev. |  | 3.603% | 2.752% | 3.265% | 2.212% | 1.851% | 8.812% | 6.592%
| PDM3 |  | -5.21% | -1.903% | 1.403% | -2.344% | -17.22% | 9.009% | 3.608%
| Z-Score (Absolute) |  | 0.45 | 0.03 | 0.51 | 0.03 | 2.2 | 1.62 | 0.83 | 4.62
| Analytical results for Fe in OREAS 906b (Certified Value 4.92 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | 2
3
| 1 | 1 | 4.9 | 4.97 | 4.8511 | 4.75 | 4.36 | 4.973 | 4.917 | 4.91 | 4.567 | 4.882 | 5.14 | 4.68 | 5.4537 | 5.323 | 4.49 | 5.03 | 5.283 | 5.02 | 5.3009 | 4.79 | 1
| 1 | 2 | 4.95 | 4.87 | 4.8397 | 4.72 | 4.36 | 4.847 | 4.931 | 5.295 | 4.791 | 4.882 | 5.14 | 4.62 | 5.4541 | 5.05 | 4.76 | 5.04 | 5.376 | 5.04 | 5.41 | 4.77 | #N/A
| 1 | 3 | 5.01 | 4.91 | 4.648 | 4.91 | 4.36 | 4.819 | 4.938 | 4.763 | 4.861 | 4.973 | 5.04 | 4.47 | 5.4596 | 5.197 | 4.57 | 5.04 | 5.27 | 5.02 | 5.2322 | 4.8 | 16
| 1 | 4 | 5.06 | 4.89 | 4.8062 | 4.97 | 4.35 | 4.91 | 5.05 | 4.826 | 4.267 | 4.791 | 4.71 | 4.56 | 5.451 | 5.302 | 4.65 | 5.03 | 5.256 | 5.03 | 5.1462 | 4.79 | 4.9154
| 1 | 5 | 4.99 | 4.85 | 4.8699 | 4.91 | 4.35 | 4.889 | 4.973 | 5.309 | 4.33 | 4.819 | 4.9 | 4.42 | 5.4556 | 5.225 | 4.76 | 5.04 | 4.994 | 5.04 | 5.1454 | 4.7 | 91
| 1 | 6 | 4.9 | 4.96 | 4.6306 | 4.83 | 4.35 | 4.931 | 4.952 | 4.98 | 4.386 | 4.819 | 4.76 | 4.49 | 5.459 | 5.134 | 5.81 | 5.04 | 5.206 | 5.05 | 5.2744 | 4.8
| Mean |  | 4.9683 | 4.9083 | 4.7742 | 4.8483 | 4.355 | 4.8948 | 4.9602 | 5.0138 | 4.5337 | 4.861 | 4.9483 | 4.54 | 5.4555 | 5.2052 | 4.84 | 5.0367 | 5.2308 | 5.0333 | 5.2515 | 4.775
| Median |  | 4.97 | 4.9 | 4.823 | 4.87 | 4.355 | 4.8995 | 4.945 | 4.945 | 4.4765 | 4.8505 | 4.97 | 4.525 | 5.4548 | 5.211 | 4.705 | 5.04 | 5.263 | 5.035 | 5.2533 | 4.79
| Std Dev. |  | 0.063692 | 0.048339 | 0.1067 | 0.098877 | 0.0054772 | 0.056073 | 0.047973 | 0.23508 | 0.24855 | 0.066112 | 0.18787 | 0.098184 | 0.003295 | 0.10282 | 0.48687 | 0.005164 | 0.12858 | 0.012111 | 0.1008 | 0.038341
| Rel.Std.Dev. |  | 1.282% | 0.9848% | 2.235% | 2.039% | 0.1258% | 1.146% | 0.9672% | 4.689% | 5.482% | 1.36% | 3.797% | 2.163% | 0.0604% | 1.975% | 10.06% | 0.1025% | 2.458% | 0.2406% | 1.919% | 0.8029%
| PDM3 |  | 1.077% | -0.1437% | -2.872% | -1.364% | -11.4% | -0.4184% | 0.9108% | 2.003% | -7.766% | -1.107% | 0.67% | -7.637% | 10.99% | 5.895% | -1.534% | 2.467% | 6.417% | 2.399% | 6.838% | -2.856%
| Z-Score (Absolute) |  | 0.25 | 0.13 | 0.97 | 0.51 | 3.62 | 0.21 | 0.2 | 0.54 | 2.5 | 0.42 | 0.13 | 2.45 | 3.33 | 1.75 | 0.56 | 0.68 | 1.91 | 0.66 | 2.04 | 0.97
| Analytical results for Ga in OREAS 906b (Certified Value 28 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.25 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [10 empty] | 1
2
| 1 | 1 | 28 | 29 | 27.2697 | 25.8 | 29 | 28 | 47.43 | 26 | 31 | 29.2 | [10 empty] | 1
| 1 | 2 | 28 | 28 | 27.0352 | 25.8 | 29 | 27 | 46.68 | 27 | 31 | 27.4 | [10 empty] | 47
| 1 | 3 | 27 | 28 | 25.6845 | 25.5 | 29 | 26 | 47.47 | 27 | 32 | 30.3 | [10 empty] | 16
| 1 | 4 | 28 | 29 | 28.372 | 24.9 | 28 | 26 | 45.79 | 27 | 32 | 26.2 | [10 empty] | 28.044
| 1 | 5 | 28 | 29 | 27.6114 | 25.9 | 29 | 28 | 46.2 | 28 | 31 | 29.3 | [10 empty] | 92
| 1 | 6 | 27 | 28 | 25.8012 | 25.4 | 29 | 27 | 46.41 | 27 | 33 | 32.9
| Mean |  | 27.6667 | 28.5 | 26.9624 | 25.55 | 28.8333 | 27 | 46.6633 | 27 | 31.6667 | 29.2167
| Median |  | 28 | 28.5 | 27.1525 | 25.65 | 29 | 27 | 46.545 | 27 | 31.5 | 29.25
| Std Dev. |  | 0.5164 | 0.54772 | 1.0478 | 0.37283 | 0.40825 | 0.89443 | 0.67533 | 0.63246 | 0.8165 | 2.3302
| Rel.Std.Dev. |  | 1.866% | 1.922% | 3.886% | 1.459% | 1.416% | 3.313% | 1.447% | 2.342% | 2.578% | 7.975%
| PDM3 |  | -1.345% | 1.626% | -3.857% | -8.893% | 2.815% | -3.723% | 66.39% | -3.723% | 12.92% | 4.182%
| Z-Score (Absolute) |  | 0.25 | 0.25 | 0.69 | 1.55 | 0.46 | 0.66 | 11.37 | 0.66 | 2.19 | 0.69
| Analytical results for Gd in OREAS 906b (Certified Value 6.05 ppm) | [21 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 5.9 | 6.2 | 6.17 | 5.58 | 5.15 | 6.56 | 7.4 | 7 | [12 empty] | 1
| 1 | 2 | 5.9 | 6.2 | 6.43 | 6 | 5.12 | 6.44 | 5.6 | 7 | [12 empty] | 19
| 1 | 3 | 5.8 | 5.5 | 6.47 | 5.64 | 5.12 | 6.45 | 7.5 | 7 | [12 empty] | 16
| 1 | 4 | 6.1 | 6 | 5.93 | 5.85 | 5.09 | 6.24 | 6.5 | 7 | [12 empty] | 6.0505
| 1 | 5 | 5.7 | 6.4 | 6.39 | 5.65 | 5.13 | 6.79 | 7.1 | 6 | [12 empty] | 93
| 1 | 6 | 5.8 | 6.2 | 6.19 | 5.79 | 5.08 | 6.06 | 6.3 | 7
| Mean |  | 5.8667 | 6.0833 | 6.2633 | 5.7517 | 5.115 | 6.4233 | 6.7333 | 6.8333
| Median |  | 5.85 | 6.2 | 6.29 | 5.72 | 5.12 | 6.445 | 6.8 | 7
| Std Dev. |  | 0.13663 | 0.31252 | 0.20578 | 0.15817 | 0.025884 | 0.25288 | 0.73394 | 0.40825
| Rel.Std.Dev. |  | 2.329% | 5.137% | 3.286% | 2.75% | 0.506% | 3.937% | 10.9% | 5.974%
| PDM3 |  | -3.038% | 0.5431% | 3.518% | -4.939% | -15.46% | 6.162% | 11.29% | 12.94%
| Z-Score (Absolute) |  | 0.44 | 0 | 0.37 | 0.67 | 1.97 | 0.69 | 1.32 | NA
| Analytical results for Ge in OREAS 906b (Certified Value 1.74 ppm) | [21 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.03 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [13 empty] | 2
2
| 1 | 1 | 2 | 1.4943 | 2 | 1 | 11.66 | < 5 | 1.8 | [13 empty] | 1
| 1 | 2 | 2 | 1.4665 | 2 | 2 | 11.4 | < 5 | 1.5 | [13 empty] | 49
| 1 | 3 | 1 | 1.0694 | 2 | 2 | 10.59 | < 5 | 1.8 | [13 empty] | 16
| 1 | 4 | 2 | 1.3505 | 2 | 2 | 10.06 | < 5 | 1.7 | [13 empty] | 1.7396
| 1 | 5 | 2 | 1.619 | 2 | 2 | 10.75 | < 5 | 1.6 | [13 empty] | 94
| 1 | 6 | 2 | 1.187 | 2 | 2 | 10.42 | < 5 | 1.6
| Mean |  | 1.8333 | 1.3645 | 2 | 1.8333 | 10.8133 |  | 1.6667
| Median |  | 2 | 1.4085 | 2 | 2 | 10.67 |  | 1.65
| Std Dev. |  | 0.40825 | 0.20533 | 0 | 0.40825 | 0.60616 |  | 0.12111
| Rel.Std.Dev. |  | 22.27% | 15.05% | 0% | 22.27% | 5.606% |  | 7.266%
| PDM3 |  | 5.391% | -21.56% | 14.97% | 5.391% | 521.6% |  | -4.19%
| Z-Score (Absolute) |  | 0 | 1.9 | 0.67 | 0 | 36.33 | 2.7 | 0.67
| Analytical results for Hf in OREAS 906b (Certified Value 6.91 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [13 empty] | 2
3
| 1 | 1 | 6.9 | 7 | 7 | 6 | 6.7 | < 1 | 10 | [13 empty] | 1
| 1 | 2 | 7 | 7.5 | 7 | 7 | 6.7 | < 1 | < 10 | [13 empty] | 33
| 1 | 3 | 8 | 7.3 | 7 | 7 | 6.6 | < 1 | 10 | [13 empty] | 16
| 1 | 4 | 7.9 | 7.1 | 7 | 6 | 6.7 | < 1 | < 10 | [13 empty] | 6.9067
| 1 | 5 | 7 | 7.3 | 7 | 6 | 6.5 | < 1 | 20 | [13 empty] | 95
| 1 | 6 | 7.1 | 7.2 | 7 | 6 | 6.7 | < 1 | < 10
| Mean |  | 7.3167 | 7.2333 | 7 | 6.3333 | 6.65 |  | 13.3333
| Median |  | 7.05 | 7.25 | 7 | 6 | 6.7 |  | 10
| Std Dev. |  | 0.49565 | 0.17512 | 0 | 0.5164 | 0.083666 |  | 5.7735
| Rel.Std.Dev. |  | 6.774% | 2.421% | 0% | 8.154% | 1.258% |  | 43.3%
| PDM3 |  | 5.936% | 4.73% | 1.351% | -8.301% | -3.716% |  | 93.05%
| Z-Score (Absolute) |  | 0.61 | 0.45 | 0 | 1.28 | 0.67 | 12.52 | 4.17
| Analytical results for Hg in OREAS 906b (Indicative Value < 0.1 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 2.06 | 2.20 | [18 empty] | ppm
|  |  | PF*OES | PF*OES | [18 empty] | 2
2
| 1 | 1 | < 5 | < 0.1 | [18 empty] | 1
| 1 | 2 | < 5 | < 0.1 | [18 empty] | 14
| 1 | 3 | < 5 | < 0.1 | [18 empty] | 16
| 1 | 4 | < 5 | < 0.1 | [18 empty] | < 0.1
| 1 | 5 | < 5 | < 0.1 | [18 empty] | 20
| 1 | 6 | < 5 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ho in OREAS 906b (Certified Value 0.52 ppm) | [21 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 0.5 | 0.5 | 0.55 | 0.51 | 0.48 | 0.49 | 0.5 | < 2 | [12 empty] | 1
| 1 | 2 | 0.5 | 0.5 | 0.57 | 0.5 | 0.49 | 0.6 | 0.4 | < 2 | [12 empty] | 16
| 1 | 3 | 0.5 | < 0.5 | 0.6 | 0.51 | 0.5 | 0.52 | 0.6 | < 2 | [12 empty] | 16
| 1 | 4 | 0.5 | 0.5 | 0.53 | 0.47 | 0.48 | 0.58 | 0.5 | < 2 | [12 empty] | 0.52095
| 1 | 5 | 0.5 | 0.5 | 0.56 | 0.51 | 0.48 | 0.58 | 0.6 | < 2 | [12 empty] | 96
| 1 | 6 | 0.6 | 0.5 | 0.54 | 0.49 | 0.5 | 0.54 | 0.6 | < 2
| Mean |  | 0.51667 | 0.5 | 0.55833 | 0.49833 | 0.48833 | 0.55167 | 0.53333
| Median |  | 0.5 | 0.5 | 0.555 | 0.505 | 0.485 | 0.56 | 0.55
| Std Dev. |  | 0.040825 | 0 | 0.024833 | 0.016021 | 0.0098319 | 0.042151 | 0.08165
| Rel.Std.Dev. |  | 7.902% | 0% | 4.448% | 3.215% | 2.013% | 7.641% | 15.31%
| PDM3 |  | -0.8227% | -4.022% | 7.176% | -4.342% | -6.261% | 5.896% | 2.377%
| Z-Score (Absolute) |  | 0.19 | 1.5 | 0.75 | 0.6 | 0.82 | 0.6 | 0.19 | 10.68
| Analytical results for In in OREAS 906b (Certified Value 1.16 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 1.2 | 1.2 | 1.1337 | 1.2 | 1.1 | 0.79 | 1.28 | 1 | [12 empty] | 1
| 1 | 2 | 1.1 | 1.2 | 1.1479 | 1.2 | 1.1 | 0.774 | 1.36 | 1.2 | [12 empty] | 35
| 1 | 3 | 1 | 1.3 | 1.0797 | 1.2 | 1 | 0.794 | 1.15 | 1.3 | [12 empty] | 16
| 1 | 4 | 1.1 | 1.3 | 1.1472 | 1.2 | 1.1 | 0.81 | 1.18 | 1.3 | [12 empty] | 1.1642
| 1 | 5 | 1.4 | 1.3 | 1.0943 | 1.2 | 1 | 0.794 | 1.14 | 1.2 | [12 empty] | 97
| 1 | 6 | 1.1 | 1.3 | 1.0856 | 1.2 | 1.1 | 0.819 | 1.1 | 1.2
| Mean |  | 1.15 | 1.2667 | 1.1147 | 1.2 | 1.0667 | 0.79683 | 1.2017 | 1.2
| Median |  | 1.1 | 1.3 | 1.114 | 1.2 | 1.1 | 0.794 | 1.165 | 1.2
| Std Dev. |  | 0.13784 | 0.05164 | 0.031653 | 0 | 0.05164 | 0.015804 | 0.098472 | 0.10954
| Rel.Std.Dev. |  | 11.99% | 4.077% | 2.839% | 0% | 4.841% | 1.983% | 8.195% | 9.129%
| PDM3 |  | -1.224% | 8.797% | -4.252% | 3.071% | -8.382% | -31.56% | 3.214% | 3.071%
| Z-Score (Absolute) |  | 0.39 | 1.42 | 0.93 | 0.39 | 1.68 | 5.87 | 0.41 | 0.39
| Analytical results for K in OREAS 906b (Certified Value 2.98 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | 2
3
| 1 | 1 | 2.86 | 3.06 | 2.4374 | 3.36 | 1.34 | 3.022 | 2.839 | 2.897 | 2.59 | 3.013 | 3 | 3.36 | 2.9242 | 3.146 | 2.95 | 3.02 | 2.649 | 3 | 3.1525 | 2.9 | 1
| 1 | 2 | 2.88 | 3.06 | 2.4966 | 3.26 | 1.34 | 2.989 | 2.839 | 2.823 | 2.731 | 3.03 | 3.1 | 3.43 | 2.936 | 3.072 | 2.95 | 3.01 | 2.649 | 2.9 | 3.216 | 2.99 | #N/A
| 1 | 3 | 2.91 | 3.02 | 2.5314 | 3.49 | 1.34 | 2.947 | 2.823 | 2.823 | 2.972 | 3.121 | 2.9 | 3.18 | 2.934 | 3.022 | 2.93 | 3.01 | 2.694 | 3 | 3.142 | 2.92 | 16
| 1 | 4 | 2.94 | 3.1 | 2.4122 | 3.29 | 1.34 | 3.005 | 2.881 | 2.789 | 2.407 | 2.989 | 3.3 | 3.28 | 2.9389 | 3.13 | 2.86 | 3.03 | 2.675 | 3 | 3.1036 | 2.95 | 2.9834
| 1 | 5 | 2.92 | 3.12 | 2.458 | 3.39 | 1.34 | 2.997 | 2.856 | 2.889 | 2.507 | 3.113 | 3.1 | 3.1 | 2.9695 | 3.138 | 2.85 | 3.02 | 2.619 | 2.9 | 3.1083 | 2.89 | 98
| 1 | 6 | 2.86 | 3.07 | 2.494 | 3.42 | 1.33 | 3.022 | 2.872 | 2.93 | 2.524 | 2.997 | 3.3 | 3.14 | 2.934 | 3.038 | 2.86 | 3.03 | 2.643 | 2.8 | 3.1884 | 2.98
| Mean |  | 2.895 | 3.0717 | 2.4716 | 3.3683 | 1.3383 | 2.997 | 2.8517 | 2.8585 | 2.6218 | 3.0438 | 3.1167 | 3.2483 | 2.9394 | 3.091 | 2.9 | 3.02 | 2.6548 | 2.9333 | 3.1518 | 2.9383
| Median |  | 2.895 | 3.065 | 2.476 | 3.375 | 1.34 | 3.001 | 2.8475 | 2.856 | 2.557 | 3.0215 | 3.1 | 3.23 | 2.935 | 3.101 | 2.895 | 3.02 | 2.649 | 2.95 | 3.1472 | 2.935
| Std Dev. |  | 0.033317 | 0.034881 | 0.043776 | 0.084715 | 0.0040825 | 0.02785 | 0.02207 | 0.054544 | 0.20222 | 0.058448 | 0.16021 | 0.13029 | 0.015558 | 0.054196 | 0.048166 | 0.0089443 | 0.026202 | 0.08165 | 0.044211 | 0.041673
| Rel.Std.Dev. |  | 1.151% | 1.136% | 1.771% | 2.515% | 0.305% | 0.9292% | 0.7739% | 1.908% | 7.713% | 1.92% | 5.14% | 4.011% | 0.5293% | 1.753% | 1.661% | 0.2962% | 0.987% | 2.784% | 1.403% | 1.418%
| PDM3 |  | -2.963% | 2.958% | -17.16% | 12.9% | -55.14% | 0.4556% | -4.416% | -4.187% | -12.12% | 2.025% | 4.467% | 8.88% | -1.474% | 3.606% | -2.796% | 1.227% | -11.01% | -1.678% | 5.644% | -1.511%
| Z-Score (Absolute) |  | 0.25 | 0.75 | 2.65 | 2.44 | 9.08 | 0.33 | 0.5 | 0.46 | 1.8 | 0.6 | 1.01 | 1.76 | 0 | 0.86 | 0.22 | 0.46 | 1.61 | 0.03 | 1.21 | 0
| Analytical results for La in OREAS 906b (Certified Value 44.2 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | [9 empty] | 1
2
| 1 | 1 | 44.5 | 43 | 46 | 46.4 | 43.9 | 33.5 | 53.5 | 44 | 24 | 46.2 | 56 | [9 empty] | 1
| 1 | 2 | 43.9 | 42 | 40.7 | 48.8 | 43 | 35 | 48.9 | 43 | 24 | 44.2 | 52 | [9 empty] | 36
| 1 | 3 | 44 | 37 | 34.9 | 47 | 43.7 | 34.1 | 44.4 | 43 | 24 | 45.8 | 51 | [9 empty] | 16
| 1 | 4 | 43.8 | 43 | 42.3 | 45.8 | 44 | 34.4 | 43.1 | 44 | 25 | 46.3 | 51 | [9 empty] | 44.2188
| 1 | 5 | 43.5 | 42 | 44.5 | 47 | 43 | 34.3 | 42.9 | 45 | 24 | 47.4 | 52 | [9 empty] | 99
| 1 | 6 | 44 | 42 | 41.9 | 47.3 | 42.7 | 34.8 | 40.8 | 43 | 23 | 47.3 | 54
| Mean |  | 43.95 | 41.5 | 41.7167 | 47.05 | 43.3833 | 34.35 | 45.6 | 43.6667 | 24 | 46.2 | 52.6667
| Median |  | 43.95 | 42 | 42.1 | 47 | 43.35 | 34.35 | 43.75 | 43.5 | 24 | 46.25 | 52
| Std Dev. |  | 0.32711 | 2.2583 | 3.8463 | 1.0114 | 0.54924 | 0.53198 | 4.7185 | 0.8165 | 0.63246 | 1.1679 | 1.9664
| Rel.Std.Dev. |  | 0.7443% | 5.442% | 9.22% | 2.15% | 1.266% | 1.549% | 10.35% | 1.87% | 2.635% | 2.528% | 3.734%
| PDM3 |  | -0.6078% | -6.148% | -5.658% | 6.403% | -1.889% | -22.32% | 3.124% | -1.249% | -45.72% | 4.481% | 19.1%
| Z-Score (Absolute) |  | 0.09 | 0.67 | 0.61 | 1.05 | 0.09 | 2.9 | 0.6 | 0 | 6.12 | 0.79 | 2.8
| Analytical results for Li in OREAS 906b (Certified Value 20.8 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.12 | 2.14 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | [8 empty] | 1
1
| 1 | 1 | 20 | 22 | 29.8694 | 24 | 15 | 18 | 18.8 | 22 | < 10 | 21 | 32 | 20 | [8 empty] | 1
| 1 | 2 | 21 | 21 | 30.258 | 23 | 22 | 17 | 19.1 | 23 | < 10 | 20 | 32 | 20 | [8 empty] | #N/A
| 1 | 3 | 20 | 21 | 27.9595 | 22 | 15 | 17 | 18.6 | 23 | < 10 | 21 | 32 | 20 | [8 empty] | 16
| 1 | 4 | 21 | 21 | 28.7096 | 26 | 29 | 18 | 22.1 | 23 | < 10 | 20 | 24 | 20 | [8 empty] | 20.8233
| 1 | 5 | 21 | 21 | 26.3129 | 22 | 21 | 17 | 18.5 | 23 | < 10 | 22 | 32 | 20 | [8 empty] | 100
| 1 | 6 | 20 | 22 | 25.2411 | 23 | 28 | 17 | 19.6 | 23 | < 10 | 20 | 18 | 20
| Mean |  | 20.5 | 21.3333 | 28.0584 | 23.3333 | 21.6667 | 17.3333 | 19.45 | 22.8333 |  | 20.6667 | 28.3333 | 20
| Median |  | 20.5 | 21 | 28.3345 | 23 | 21.5 | 17 | 18.95 | 23 |  | 20.5 | 32 | 20
| Std Dev. |  | 0.54772 | 0.5164 | 1.9766 | 1.5055 | 6.0553 | 0.5164 | 1.3576 | 0.40825 |  | 0.8165 | 5.9889 | 0
| Rel.Std.Dev. |  | 2.672% | 2.421% | 7.045% | 6.452% | 27.95% | 2.979% | 6.98% | 1.788% |  | 3.951% | 21.14% | 0%
| PDM3 |  | -1.553% | 2.449% | 34.75% | 12.05% | 4.05% | -16.76% | -6.595% | 9.653% |  | -0.7524% | 36.07% | -3.954%
| Z-Score (Absolute) |  | 0.3 | 0 | 2.41 | 0.72 | 0.12 | 1.43 | 0.67 | 0.54 | 5.85 | 0.24 | 2.51 | NA
Results from laboratory 2.27 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 906b (Indicative Value 0.069 ppm) | [21 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.26 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | [14 empty] | 3
3
| 1 | 1 | < 0.5 | 0.08 | < 0.05 | 0.07 | 0.08 | NA | [14 empty] | 1
| 1 | 2 | < 0.5 | 0.08 | 0.05 | 0.07 | 0.08 | NA | [14 empty] | 9
| 1 | 3 | < 0.5 | 0.1 | < 0.05 | 0.08 | 0.08 | NA | [14 empty] | 16
| 1 | 4 | < 0.5 | 0.08 | 0.05 | 0.07 | 0.08 | NA | [14 empty] | 0.068542
| 1 | 5 | < 0.5 | 0.08 | 0.05 | 0.07 | 0.07 | NA | [14 empty] | 15
| 1 | 6 | < 0.5 | 0.08 | < 0.05 | 0.08 | 0.09 | NA
| Mean |  |  | 0.083333 | 0.05 | 0.073333 | 0.08
| Median |  |  | 0.08 | 0.05 | 0.07 | 0.08
| Std Dev. |  |  | 0.008165 | 8.4984e-18 | 0.005164 | 0.0063246
| Rel.Std.Dev. |  |  | 9.798% | 1.7e-14% | 7.042% | 7.906%
| PDM3 |  |  | 21.58% | -27.05% | 6.991% | 16.72%
| Z-Score (Absolute) |  | 17.19 | 0.34 | 4.3 | 0.67 | 0 | NA
| Analytical results for Mg in OREAS 906b (Certified Value 0.233 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.24 | 0.23 | 0.23999 | 0.23 | 0.21 | 0.241 | 0.229 | 0.211 | 0.229 | 0.241 | 0.24 | 0.25 | 0.23555 | 0.241 | 0.23 | 0.34 | 0.231 | 0.24 | 0.23 |  | 1
| 1 | 2 | 0.24 | 0.24 | 0.23062 | 0.22 | 0.21 | 0.235 | 0.229 | 0.271 | 0.241 | 0.241 | 0.26 | 0.26 | 0.23894 | 0.241 | 0.24 | 0.33 | 0.24 | 0.23 | 0.24 |  | #N/A
| 1 | 3 | 0.25 | 0.25 | 0.20937 | 0.23 | 0.21 | 0.241 | 0.229 | 0.211 | 0.229 | 0.241 | 0.24 | 0.24 | 0.23809 | 0.235 | 0.23 | 0.33 | 0.231 | 0.23 | 0.22 |  | 16
| 1 | 4 | 0.25 | 0.24 | 0.22596 | 0.23 | 0.22 | 0.241 | 0.241 | 0.217 | 0.217 | 0.241 | 0.22 | 0.25 | 0.23252 | 0.241 | 0.24 | 0.33 | 0.229 | 0.23 | 0.23 |  | 0.23285
| 1 | 5 | 0.24 | 0.24 | 0.24073 | 0.23 | 0.21 | 0.241 | 0.229 | 0.211 | 0.217 | 0.235 | 0.23 | 0.26 | 0.2315 | 0.241 | 0.23 | 0.33 | 0.225 | 0.24 | 0.22 |  | 101
| 1 | 6 | 0.24 | 0.24 | 0.21426 | 0.22 | 0.21 | 0.241 | 0.229 | 0.217 | 0.217 | 0.241 | 0.22 | 0.25 | 0.23898 | 0.253 | 0.24 | 0.34 | 0.231 | 0.24 | 0.23
| Mean |  | 0.24333 | 0.24 | 0.22682 | 0.22667 | 0.21167 | 0.24 | 0.231 | 0.223 | 0.225 | 0.24 | 0.235 | 0.25167 | 0.23593 | 0.242 | 0.235 | 0.33333 | 0.23117 | 0.235 | 0.22833
| Median |  | 0.24 | 0.24 | 0.22829 | 0.23 | 0.21 | 0.241 | 0.229 | 0.214 | 0.223 | 0.241 | 0.235 | 0.25 | 0.23682 | 0.241 | 0.235 | 0.33 | 0.231 | 0.235 | 0.23
| Std Dev. |  | 0.005164 | 0.0063246 | 0.012996 | 0.005164 | 0.0040825 | 0.0024495 | 0.004899 | 0.023698 | 0.009798 | 0.0024495 | 0.015166 | 0.0075277 | 0.0033007 | 0.0058992 | 0.0054772 | 0.005164 | 0.004916 | 0.0054772 | 0.0075277
| Rel.Std.Dev. |  | 2.122% | 2.635% | 5.73% | 2.278% | 1.929% | 1.021% | 2.121% | 10.63% | 4.355% | 1.021% | 6.454% | 2.991% | 1.399% | 2.438% | 2.331% | 1.549% | 2.127% | 2.331% | 3.297%
| PDM3 |  | 4.502% | 3.071% | -2.589% | -2.655% | -9.097% | 3.071% | -0.7943% | -4.23% | -3.371% | 3.071% | 0.9235% | 8.081% | 1.323% | 3.93% | 0.9235% | 43.15% | -0.7228% | 0.9235% | -1.94%
| Z-Score (Absolute) |  | 0.84 | 0.51 | 0.83 | 0.84 | 2.36 | 0.53 | 0.39 | 1.2 | 1 | 0.53 | 0 | 1.69 | 0.09 | 0.73 | 0 | 9.95 | 0.39 | 0 | 0.67
| Analytical results for Mn in OREAS 906b (Certified Value 0.032 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 3
3
| 1 | 1 | 0.031 | 0.033485 | 0.03 | 0.03495 | 0.031 | 0.031 | 0.031 | 0.031 | 0.031 | 0.0336 | 0.0324 | 0.035461 | 0.031 | 0.031 | 0.0377 | 0.0328 | 0.0351 | 0.03 |  |  | 1
| 1 | 2 | 0.031 | 0.031484 | 0.03 | 0.03391 | 0.031 | 0.031 | 0.039 | 0.039 | 0.031 | 0.0342 | 0.0328 | 0.035956 | 0.031 | 0.0318 | 0.0378 | 0.0322 | 0.0349 | 0.03 |  |  | 38
| 1 | 3 | 0.03 | 0.030381 | 0.03 | 0.03576 | 0.031 | 0.031 | 0.031 | 0.023 | 0.031 | 0.0342 | 0.0332 | 0.035797 | 0.031 | 0.0302 | 0.0378 | 0.0324 | 0.0351 | 0.03 |  |  | 16
| 1 | 4 | 0.031 | 0.031723 | 0.03 | 0.03383 | 0.031 | 0.031 | 0.031 | 0.031 | 0.031 | 0.032 | 0.0333 | 0.0355 | 0.031 | 0.031 | 0.0377 | 0.0322 | 0.0337 | 0.03 |  |  | 0.032176
| 1 | 5 | 0.032 | 0.033177 | 0.03 | 0.0332 | 0.031 | 0.031 | 0.031 | 0.031 | 0.031 | 0.0319 | 0.0327 | 0.035204 | 0.039 | 0.0318 | 0.0378 | 0.0328 | 0.033 | 0.03 |  |  | 102
| 1 | 6 | 0.031 | 0.031228 | 0.03 | 0.03514 | 0.031 | 0.031 | 0.031 | 0.031 | 0.031 | 0.0317 | 0.032 | 0.035229 | 0.031 | 0.0457 | 0.0378 | 0.033 | 0.0348 | 0.03
| Mean |  | 0.031 | 0.031913 | 0.03 | 0.034465 | 0.031 | 0.031 | 0.032333 | 0.031 | 0.031 | 0.032933 | 0.032733 | 0.035525 | 0.032333 | 0.033583 | 0.037767 | 0.032567 | 0.034433 | 0.03
| Median |  | 0.031 | 0.031603 | 0.03 | 0.03443 | 0.031 | 0.031 | 0.031 | 0.031 | 0.031 | 0.0328 | 0.03275 | 0.03548 | 0.031 | 0.0314 | 0.0378 | 0.0326 | 0.03485 | 0.03
| Std Dev. |  | 0.00063246 | 0.0011921 | 0 | 0.0009674 | 0 | 0 | 0.003266 | 0.0050596 | 0 | 0.0011928 | 0.00048854 | 0.00030187 | 0.003266 | 0.005966 | 5.164e-05 | 0.00034448 | 0.0008756 | 0
| Rel.Std.Dev. |  | 2.04% | 3.735% | 0% | 2.807% | 0% | 0% | 10.1% | 16.32% | 0% | 3.622% | 1.492% | 0.8497% | 10.1% | 17.76% | 0.1367% | 1.058% | 2.543% | 0%
| PDM3 |  | -3.654% | -0.817% | -6.762% | 7.115% | -3.654% | -3.654% | 0.4895% | -3.654% | -3.654% | 2.354% | 1.733% | 10.41% | 0.4895% | 4.374% | 17.38% | 1.215% | 7.016% | -6.762%
| Z-Score (Absolute) |  | 0.66 | 0.19 | 1.19 | 1.15 | 0.67 | 0.67 | 0 | 0.67 | 0.67 | 0.35 | 0.24 | 1.7 | 0 | 0.67 | 2.87 | 0.16 | 1.13 | 1.19
| Analytical results for Mo in OREAS 906b (Certified Value 5 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 5 | 5 | 4.7017 | 3.2 | 5 | 5 | 3.63 | 4 | 6 | < 10 | 5 | [9 empty] | 1
| 1 | 2 | 5 | 5 | 4.7121 | 3 | 6 | 5 | 3.59 | 5 | 6 | < 10 | 6 | [9 empty] | 39
| 1 | 3 | 5 | 5 | 3.919 | 3.1 | 5 | 6 | 3.55 | 5 | 6 | < 10 | 5 | [9 empty] | 16
| 1 | 4 | 5 | 55 | 4.8883 | 3.1 | 5 | 5 | 3.52 | 4 | 6 | < 10 | 6 | [9 empty] | 4.9994
| 1 | 5 | 5 | 5 | 4.3752 | 3.1 | 5 | 6 | 3.43 | 4 | 6 | < 10 | 6 | [9 empty] | 103
| 1 | 6 | 5 | 5 | 4.153 | 3 | 5 | 5 | 3.5 | 6 | 6 | < 10 | 7
| Mean |  | 5 | 13.3333 | 4.4582 | 3.0833 | 5.1667 | 5.3333 | 3.5367 | 4.6667 | 6 |  | 5.8333
| Median |  | 5 | 5 | 4.5384 | 3.1 | 5 | 5 | 3.535 | 4.5 | 6 |  | 6
| Std Dev. |  | 0 | 20.4124 | 0.37401 | 0.075277 | 0.40825 | 0.5164 | 0.070333 | 0.8165 | 0 |  | 0.75277
| Rel.Std.Dev. |  | 0% | 153.1% | 8.389% | 2.441% | 7.902% | 9.682% | 1.989% | 17.5% | 0% |  | 12.9%
| PDM3 |  | 0.01139% | 166.7% | -10.83% | -38.33% | 3.345% | 6.679% | -29.26% | -6.656% | 20.01% |  | 16.68%
| Z-Score (Absolute) |  | 0 | 10.37 | 0.67 | 2.39 | 0.21 | 0.41 | 1.82 | 0.41 | 1.24 | 0 | 1.04
| Analytical results for Na in OREAS 906b (Indicative Value 0.833 wt.%) | [21 empty] | Indicative
Na | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 2.06 | 2.19 | [18 empty] | wt.%
|  |  | PF*OES | PF*OES/MS | [18 empty] | 3
3
| 1 | 1 | 0.08 | 1.56 | [18 empty] | 1
| 1 | 2 | 0.09 | 1.62 | [18 empty] | 10
| 1 | 3 | 0.08 | 1.57 | [18 empty] | 16
| 1 | 4 | 0.09 | 1.6 | [18 empty] | 0.8325
| 1 | 5 | 0.08 | 1.57 | [18 empty] | 16
| 1 | 6 | 0.08 | 1.57
| Mean |  | 0.083333 | 1.5817
| Median |  | 0.08 | 1.57
| Std Dev. |  | 0.005164 | 0.023166
| Rel.Std.Dev. |  | 6.197% | 1.465%
| PDM3 |  | -89.99% | 89.99%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 906b (Certified Value 17.7 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [11 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [11 empty] | 1
1
| 1 | 1 | 25 | 15.6387 | 20 | 18.8 | 17 | 18 | 18.8 | 23 | 18.8 | [11 empty] | 1
| 1 | 2 | 30 | 16.4446 | 16 | 20.1 | 17 | 17 | 19.2 | 20 | 16.9 | [11 empty] | 40
| 1 | 3 | 25 | 12.6545 | 12 | 19.8 | 18 | 16 | 18.8 | 22 | 19 | [11 empty] | 16
| 1 | 4 | 26 | 16.6295 | 13 | 19.7 | 18 | 15 | 19.3 | 21 | 17.4 | [11 empty] | 17.6687
| 1 | 5 | 27 | 15.2861 | 28 | 20.3 | 17 | 17 | 19.1 | 21 | 17.9 | [11 empty] | 104
| 1 | 6 | 26 | 11.9436 | 10 | 20.5 | 17 | 16 | 19.2 | 22 | 18.7
| Mean |  | 26.5 | 14.7662 | 16.5 | 19.8667 | 17.3333 | 16.5 | 19.0667 | 21.5 | 18.1167
| Median |  | 26 | 15.4624 | 14.5 | 19.95 | 17 | 16.5 | 19.15 | 21.5 | 18.3
| Std Dev. |  | 1.8708 | 1.9873 | 6.6257 | 0.60222 | 0.5164 | 1.0488 | 0.21602 | 1.0488 | 0.85186
| Rel.Std.Dev. |  | 7.06% | 13.46% | 40.16% | 3.031% | 2.979% | 6.356% | 1.133% | 4.878% | 4.702%
| PDM3 |  | 49.98% | -16.43% | -6.614% | 12.44% | -1.898% | -6.614% | 7.912% | 21.68% | 2.535%
| Z-Score (Absolute) |  | 3.5 | 1.4 | 0.67 | 0.73 | 0.33 | 0.67 | 0.4 | 1.41 | 0
| Analytical results for Nd in OREAS 906b (Certified Value 38.4 ppm) | [21 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 1
2
| 1 | 1 | 38.2 | 39.62 | 39.8 | 36.3 | 31.7 | 36.2 | 41.9 | 45 | [12 empty] | 1
| 1 | 2 | 37.5 | 36.23 | 41.6 | 35.4 | 33.1 | 42.7 | 36.9 | 42 | [12 empty] | 18
| 1 | 3 | 38.1 | 38.83 | 40.4 | 35.2 | 32.5 | 41 | 42.5 | 40 | [12 empty] | 16
| 1 | 4 | 38.2 | 36.43 | 39.1 | 36.4 | 32 | 39.5 | 38.2 | 41 | [12 empty] | 38.3598
| 1 | 5 | 37.7 | 38.51 | 40.3 | 35 | 32.1 | 38.6 | 45.5 | 41 | [12 empty] | 105
| 1 | 6 | 37.1 | 36.35 | 41 | 34.6 | 32.5 | 35.3 | 49.2 | 43
| Mean |  | 37.8 | 37.6617 | 40.3667 | 35.4833 | 32.3167 | 38.8833 | 42.3667 | 42
| Median |  | 37.9 | 37.47 | 40.35 | 35.3 | 32.3 | 39.05 | 42.2 | 41.5
| Std Dev. |  | 0.44721 | 1.4971 | 0.87788 | 0.72226 | 0.4916 | 2.8124 | 4.5606 | 1.7889
| Rel.Std.Dev. |  | 1.183% | 3.975% | 2.175% | 2.036% | 1.521% | 7.233% | 10.76% | 4.259%
| PDM3 |  | -1.459% | -1.82% | 5.232% | -7.499% | -15.75% | 1.365% | 10.45% | 9.49%
| Z-Score (Absolute) |  | 0.15 | 0.19 | 0.56 | 0.79 | 1.66 | 0.15 | 1.11 | 1.01
| Analytical results for Ni in OREAS 906b (Indicative Value 8 ppm) | [21 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 2
2
| 1 | 1 | < 20 | < 20 | 26 | 1.85 | < 20 | < 20 | 70 | < 20 | < 20 | < 5 | 8 | 12.9 | 6 | < 10 | 10 | < 50 |  |  |  |  | 1
| 1 | 2 | < 20 | < 20 | 13 | 2.12 | 20 | < 20 | 80 | < 20 | < 20 | 6 | 8 | 13 | < 5 | < 10 | 20 | < 50 |  |  |  |  | 11
| 1 | 3 | < 20 | < 20 | 16 | 2.05 | < 20 | < 20 | 70 | < 20 | < 20 | < 5 | 7 | 12.9 | < 5 | < 10 | < 10 | < 50 |  |  |  |  | 16
| 1 | 4 | < 20 | < 20 | 10 | 1.86 | < 20 | < 20 | 80 | < 20 | < 20 | < 5 | 6 | 12.4 | < 5 | < 10 | < 10 | < 50 |  |  |  |  | 8.0023
| 1 | 5 | < 20 | < 20 | 15 | 1.84 | < 20 | < 20 | 30 | < 20 | < 20 | < 5 | 7 | 13.6 | < 5 | < 10 | < 10 | < 50 |  |  |  |  | 17
| 1 | 6 | < 20 | < 20 | 12 | 2.09 | < 20 | < 20 | 80 | < 20 | < 20 | 6 | 7 | 13.3 | < 5 | < 10 | 20 | < 50
| Mean |  |  |  | 15.3333 | 1.9683 | 20 |  | 68.3333 |  |  | 6 | 7.1667 | 13.0167 | 6 |  | 16.6667
| Median |  |  |  | 14 | 1.955 | 20 |  | 75 |  |  | 6 | 7 | 12.95 | 6 |  | 20
| Std Dev. |  |  |  | 5.6451 | 0.13167 |  |  | 19.4079 |  |  | 0 | 0.75277 | 0.40702 |  |  | 5.7735
| Rel.Std.Dev. |  |  |  | 36.82% | 6.689% |  |  | 28.4% |  |  | 0% | 10.5% | 3.127% |  |  | 34.64%
| PDM3 |  |  |  | 91.61% | -75.4% | 149.9% |  | 753.9% |  |  | -25.02% | -10.44% | 62.66% | -25.02% |  | 108.3%
| Z-Score (Absolute) |  | 0 | 0 | 1.23 | 1.85 | 0.38 | 0 | 13.45 | 0 | 0 | 1.46 | 0.65 | 0.7 | 1.59 | 1.15 | 0.19 | 3.46
| Analytical results for P in OREAS 906b (Certified Value 0.029 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | [10 empty] | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | [10 empty] | 3
3
| 1 | 1 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.05 | 0.035 | 0.0532 | 0.0377 | 0.028 | [10 empty] | 1
| 1 | 2 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.05 | 0.026 | 0.0275 | 0.0377 | 0.028 | [10 empty] | #N/A
| 1 | 3 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.04 | 0.031 | 0.0153 | 0.0377 | 0.022 | [10 empty] | 16
| 1 | 4 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.05 | 0.031 | 0.0135 | 0.0374 | 0.022 | [10 empty] | 0.028886
| 1 | 5 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.04 | 0.039 | 0.0353 | 0.0374 | 0.022 | [10 empty] | 106
| 1 | 6 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.04 | 0.035 | 0.0812 | 0.0377 | 0.022
| Mean |  | 0.03 | 0.03 | 0.026667 | 0.02 | 0.03 | 0.045 | 0.032833 | 0.037667 | 0.0376 | 0.024
| Median |  | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.045 | 0.033 | 0.0314 | 0.0377 | 0.022
| Std Dev. |  | 0 | 0 | 0.005164 | 0 | 0 | 0.0054772 | 0.0044907 | 0.025806 | 0.00015492 | 0.0030984
| Rel.Std.Dev. |  | 0% | 0% | 19.36% | 0% | 0% | 12.17% | 13.68% | 68.51% | 0.412% | 12.91%
| PDM3 |  | 3.858% | 3.858% | -7.682% | -30.76% | 3.858% | 55.79% | 13.67% | 30.4% | 30.17% | -16.91%
| Z-Score (Absolute) |  | 0 | 0 | 0.48 | 1.44 | 0 | 2.17 | 0.39 | 1.11 | 1.1 | 0.87
| Analytical results for Pb in OREAS 906b (Certified Value 31.4 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  | 1
1
| 1 | 1 | 33 | 30 | 33.2073 | < 100 | 25 | < 100 | < 100 | 100 | 100 | 100 | 31 | 30 | 15.1 | 32 | 35 | 36 | < 300 |  |  |  | 1
| 1 | 2 | 34 | 30 | 33.4005 | < 100 | 23.4 | < 100 | < 100 | 100 | 100 | 100 | 31 | 29 | 16.1 | 35 | 39 | 39.4 | < 300 |  |  |  | 29
| 1 | 3 | 34 | 30 | 24.5374 | < 100 | 23.7 | < 100 | < 100 | 100 | 100 | 100 | 30 | 28 | 16.2 | 35 | 39 | 37.5 | < 300 |  |  |  | 16
| 1 | 4 | 34 | 30 | 26.1218 | < 100 | 24.7 | < 100 | < 100 | 100 | 100 | 100 | 30 | 29 | 15 | 33 | 34 | 35.2 | < 300 |  |  |  | 31.3709
| 1 | 5 | 33 | 30 | 29.1087 | < 100 | 24.7 | < 100 | < 100 | < 100 | 100 | 100 | 30 | 30 | 15.4 | 33 | 33 | 36.3 | < 300 |  |  |  | 107
| 1 | 6 | 34 | 30 | 26.8527 | < 100 | 24.7 | < 100 | < 100 | 100 | 100 | 100 | 29 | 31 | 15.3 | 30 | 37 | 35.2 | < 300
| Mean |  | 33.6667 | 30 | 28.8714 |  | 24.3667 |  |  | 100 | 100 | 100 | 30.1667 | 29.5 | 15.5167 | 33 | 36.1667 | 36.6
| Median |  | 34 | 30 | 27.9807 |  | 24.7 |  |  | 100 | 100 | 100 | 30 | 29.5 | 15.35 | 33 | 36 | 36.15
| Std Dev. |  | 0.5164 | 0 | 3.736 |  | 0.65013 |  |  | 0 | 0 | 0 | 0.75277 | 1.0488 | 0.51153 | 1.8974 | 2.5626 | 1.6137
| Rel.Std.Dev. |  | 1.534% | 0% | 12.94% |  | 2.668% |  |  | 0% | 0% | 0% | 2.495% | 3.555% | 3.297% | 5.75% | 7.085% | 4.409%
| PDM3 |  | 7.318% | -4.37% | -7.968% |  | -22.33% |  |  | 218.8% | 218.8% | 218.8% | -3.839% | -5.964% | -50.54% | 5.193% | 15.29% | 16.67%
| Z-Score (Absolute) |  | 0.14 | 0.35 | 0.42 | 0.79 | 0.67 | 0.79 | 0.79 | 3.17 | 3.65 | 3.65 | 0.34 | 0.38 | 1.18 | 0.18 | 0 | 0.02 | 6.5
| Analytical results for Pr in OREAS 906b (Certified Value 10.2 ppm) | [21 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 1
2
| 1 | 1 | 10.1 | 10.5 | 10.38 | 9.05 | 8.7 | 11.2 | 10.7 | 12 | [12 empty] | 1
| 1 | 2 | 10 | 9.7 | 10.78 | 8.92 | 8.92 | 10.9 | 9.5 | 11 | [12 empty] | 20
| 1 | 3 | 9.9 | 8.5 | 10.39 | 8.79 | 8.81 | 10.4 | 10.5 | 11 | [12 empty] | 16
| 1 | 4 | 10.1 | 9.9 | 10.09 | 8.85 | 8.84 | 10.2 | 9.2 | 11 | [12 empty] | 10.1537
| 1 | 5 | 10 | 10.3 | 10.5 | 8.91 | 8.79 | 8.73 | 11 | 11 | [12 empty] | 108
| 1 | 6 | 10.1 | 9.8 | 10.53 | 8.71 | 8.84 | 9.61 | 11.1 | 12
| Mean |  | 10.0333 | 9.7833 | 10.445 | 8.8717 | 8.8167 | 10.1733 | 10.3333 | 11.3333
| Median |  | 10.05 | 9.85 | 10.445 | 8.88 | 8.825 | 10.3 | 10.6 | 11
| Std Dev. |  | 0.08165 | 0.69976 | 0.22616 | 0.11737 | 0.072296 | 0.8983 | 0.79666 | 0.5164
| Rel.Std.Dev. |  | 0.8138% | 7.153% | 2.165% | 1.323% | 0.82% | 8.83% | 7.71% | 4.556%
| PDM3 |  | -1.185% | -3.647% | 2.869% | -12.63% | -13.17% | 0.1937% | 1.769% | 11.62%
| Z-Score (Absolute) |  | 0 | 0.56 | 0.93 | 2.61 | 2.73 | 0.31 | 0.67 | NA
Results from laboratory 2.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 906b (Certified Value 137 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [11 empty] | 0
0
| 1 | 1 | 136.2 | 138 | 136.8796 | 141 | 134.4 | 137 | 135 | 139 | 95 | [11 empty] | 1
| 1 | 2 | 135.8 | 140 | 130.0321 | 144 | 139.2 | 141.1 | 137 | 139 | 81 | [11 empty] | 45
| 1 | 3 | 137 | 141 | 129.2196 | 140 | 132 | 139.7 | 135 | 141 | 134 | [11 empty] | 16
| 1 | 4 | 136.6 | 140 | 134.6206 | 139 | 131.3 | 141.3 | 138 | 145 | 103 | [11 empty] | 137.1498
| 1 | 5 | 137.5 | 142 | 136.6543 | 141 | 131.8 | 139.4 | 131 | 138 | 68 | [11 empty] | 109
| 1 | 6 | 136.6 | 139 | 132.3836 | 141 | 132 | 140.4 | 135 | 138 | 57
| Mean |  | 136.6167 | 140 | 133.2983 | 141 | 133.45 | 139.8167 | 135.1667 | 140 | 89.6667
| Median |  | 136.6 | 140 | 133.5021 | 141 | 132 | 140.05 | 135 | 139 | 88
| Std Dev. |  | 0.5947 | 1.4142 | 3.2857 | 1.6733 | 3.0171 | 1.569 | 2.4014 | 2.6833 | 27.5076
| Rel.Std.Dev. |  | 0.4353% | 1.01% | 2.465% | 1.187% | 2.261% | 1.122% | 1.777% | 1.917% | 30.68%
| PDM3 |  | -0.3887% | 2.078% | -2.808% | 2.807% | -2.698% | 1.945% | -1.446% | 2.078% | -34.62%
| Z-Score (Absolute) |  | 0 | 0.69 | 0.67 | 0.89 | 0.64 | 0.65 | 0.29 | 0.69 | 9.54
| Analytical results for Re in OREAS 906b (Indicative Value 0.01 ppm) | [21 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.02 | 2.03 | 2.14 | 2.19 | 2.20 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | [15 empty] | 3
3
| 1 | 1 | < 0.1 | < 0.1 | 0.01 | 0.007 | < 0.1 | [15 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | 0.02 | 0.007 | < 0.1 | [15 empty] | 12
| 1 | 3 | < 0.1 | < 0.1 | 0.01 | 0.007 | < 0.1 | [15 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | 0.01 | 0.006 | < 0.1 | [15 empty] | 0.0099167
| 1 | 5 | < 0.1 | < 0.1 | 0.01 | 0.006 | < 0.1 | [15 empty] | 18
| 1 | 6 | < 0.1 | < 0.1 | 0.02 | 0.006 | < 0.1
| Mean |  |  |  | 0.013333 | 0.0065
| Median |  |  |  | 0.01 | 0.0065
| Std Dev. |  |  |  | 0.005164 | 0.00054772
| Rel.Std.Dev. |  |  |  | 38.73% | 8.427%
| PDM3 |  |  |  | 34.45% | -34.45%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 906b (Indicative Value 0.047 wt.%) | [21 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.20 | 2.22 | 2.25 | 2.27 | [7 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | [7 empty] | 3
3
| 1 | 1 | < 0.05 | 0.076116 | 0.04 | 0.04 | 0.05 | 0.05 | 0.2 | 0.06 | 0.04 | 0.03 | 0.05 | < 0.01 | 0.04 | [7 empty] | 1
| 1 | 2 | < 0.05 | 0.065737 | 0.05 | 0.05 | 0.05 | 0.05 | 0.57 | 0.07 | 0.04 | 0.04 | 0.05 | 0.03 | 0.04 | [7 empty] | 3
| 1 | 3 | < 0.05 | 0.073376 | 0.04 | 0.04 | 0.06 | 0.05 | 0.1 | 0.06 | 0.04 | 0.04 | 0.05 | 0.03 | 0.04 | [7 empty] | 16
| 1 | 4 | < 0.05 | 0.067213 | 0.03 | 0.04 | 0.06 | 0.05 | 0.17 | 0.06 | 0.04 | 0.04 | 0.05 | 0.01 | 0.05 | [7 empty] | 0.046926
| 1 | 5 | < 0.05 | 0.074423 | 0.04 | 0.04 | 0.06 | 0.05 | 0.25 | 0.08 | 0.04 | 0.04 | 0.05 | < 0.01 | 0.04 | [7 empty] | 19
| 1 | 6 | < 0.05 | 0.065274 | 0.04 | 0.04 | 0.06 | 0.05 | 0.07 | 0.07 | 0.03 | 0.04 | 0.05 | 0.02 | 0.04
| Mean |  |  | 0.070356 | 0.04 | 0.041667 | 0.056667 | 0.05 | 0.22667 | 0.066667 | 0.038333 | 0.038333 | 0.05 | 0.0225 | 0.041667
| Median |  |  | 0.070294 | 0.04 | 0.04 | 0.06 | 0.05 | 0.185 | 0.065 | 0.04 | 0.04 | 0.05 | 0.025 | 0.04
| Std Dev. |  |  | 0.0048142 | 0.0063246 | 0.0040825 | 0.005164 | 7.6012e-18 | 0.18052 | 0.008165 | 0.0040825 | 0.0040825 | 7.6012e-18 | 0.0095743 | 0.0040825
| Rel.Std.Dev. |  |  | 6.843% | 15.81% | 9.798% | 9.113% | 1.52e-14% | 79.64% | 12.25% | 10.65% | 10.65% | 1.52e-14% | 42.55% | 9.798%
| PDM3 |  |  | 49.93% | -14.76% | -11.21% | 20.76% | 6.55% | 383% | 42.07% | -18.31% | -18.31% | 6.55% | -52.05% | -11.21%
| Z-Score (Absolute) |  | 1.35 | 2.32 | 0.13 | 0 | 1.21 | 0.67 | 14.97 | 2.02 | 0.27 | 0.27 | 0.67 | 2.02 | 0
| Analytical results for Sb in OREAS 906b (Certified Value 2.05 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.12 | 2.14 | 2.20 | 2.24 | 2.25 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | 2.6 | 1.6798 | < 5 | 2 | 2 | 2 | < 50 | < 2 | [12 empty] | 1
| 1 | 2 | 2.2 | 1.4616 | < 5 | 2.1 | 2.2 | 2.1 | < 50 | 2 | [12 empty] | 47
| 1 | 3 | 1.9 | 1.4988 | < 5 | 2.1 | 2.1 | 2.2 | < 50 | 3 | [12 empty] | 16
| 1 | 4 | 2.7 | 1.8238 | < 5 | 2.1 | 2.1 | 2 | < 50 | 3 | [12 empty] | 2.045
| 1 | 5 | 2.1 | 1.8123 | < 5 | 2 | 2.2 | 2.1 | < 50 | 3 | [12 empty] | 110
| 1 | 6 | 2.4 | 1.6745 | < 5 | 2 | 2.1 | 2.1 | < 50 | 4
| Mean |  | 2.3167 | 1.6585 |  | 2.05 | 2.1167 | 2.0833 |  | 3
| Median |  | 2.3 | 1.6772 |  | 2.05 | 2.1 | 2.1 |  | 3
| Std Dev. |  | 0.30605 | 0.15226 |  | 0.054772 | 0.075277 | 0.075277 |  | 0.70711
| Rel.Std.Dev. |  | 13.21% | 9.181% |  | 2.672% | 3.556% | 3.613% |  | 23.57%
| PDM3 |  | 13.28% | -18.9% |  | 0.2431% | 3.503% | 1.873% |  | 46.7%
| Z-Score (Absolute) |  | 0.67 | 1.54 | 1.29 | 0.22 | 0 | 0.11 | 77.15 | NA
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 906b (Indicative Value 6.33 ppm) | [21 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.26 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | [10 empty] | 2
2
| 1 | 1 | < 20 | < 20 | 5.6 | < 5 | < 5 | 26.9 | 7 | 9 | < 5 | 4 | [10 empty] | 1
| 1 | 2 | < 20 | < 20 | 5.3 | < 5 | < 5 | 28 | 5 | 8 | < 5 | 4 | [10 empty] | 14
| 1 | 3 | < 20 | < 20 | 4.5 | < 5 | < 5 | 27.4 | 7 | 8 | < 5 | 4 | [10 empty] | 16
| 1 | 4 | < 20 | < 20 | 5.3 | < 5 | < 5 | 27.2 | 10 | 8 | < 5 | 4 | [10 empty] | 6.33
| 1 | 5 | < 20 | < 20 | 5.4 | < 5 | < 5 | 27.2 | 8 | 8 | < 5 | 4 | [10 empty] | 20
| 1 | 6 | < 20 | < 20 | 5 | < 5 | < 5 | 27.9 | 9 | 9 | < 5 | 4
| Mean |  |  |  | 5.1833 |  |  | 27.4333 | 7.6667 | 8.3333 |  | 4
| Median |  |  |  | 5.3 |  |  | 27.3 | 7.5 | 8 |  | 4
| Std Dev. |  |  |  | 0.38687 |  |  | 0.43205 | 1.7512 | 0.5164 |  | 0
| Rel.Std.Dev. |  |  |  | 7.464% |  |  | 1.575% | 22.84% | 6.197% |  | 0%
| PDM3 |  |  |  | -18.11% |  |  | 333.4% | 21.12% | 31.65% |  | -36.81%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.23 | 0.74 | 0.74 | 3.96 | 0.23 | 0.36 | 0.74 | 0.46
| Analytical results for Se in OREAS 906b (Indicative Value 5 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 2.03 | 2.19 | 2.20 | 2.25 | [16 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [16 empty] | 2
2
| 1 | 1 | < 20 | 3 | 6 | < 8 | [16 empty] | 1
| 1 | 2 | < 20 | 4 | 4 | < 8 | [16 empty] | 49
| 1 | 3 | < 20 | 3 | 4 | 12 | [16 empty] | 16
| 1 | 4 | < 20 | 3 | 4 | 8 | [16 empty] | 5
| 1 | 5 | < 20 | 3 | 5 | 11 | [16 empty] | 15
| 1 | 6 | < 20 | 3 | 5 | < 8
| Mean |  |  | 3.1667 | 4.6667 | 10.3333
| Median |  |  | 3 | 4.5 | 11
| Std Dev. |  |  | 0.40825 | 0.8165 | 2.0817
| Rel.Std.Dev. |  |  | 12.89% | 17.5% | 20.15%
| PDM3 |  |  | -36.67% | -6.667% | 106.7%
| Z-Score (Absolute) |  | 1.38 | 0.93 | 0.42 | 0.42
| Analytical results for Si in OREAS 906b (Certified Value 31.18 wt.%) | [21 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 30.9 | 30.5175 | 28.63 | 31.51 | 31.6 | 32.02 | 30.39 | 31.69 | > 30 | 29.8163 | 30.974 | 31.27 | 32.22 | 32.109 | > 30 | [5 empty] | 1
| 1 | 2 | 31.1 | 30.1964 | 28.53 | 31.09 | 31.41 | 31.79 | 31.88 | 31.88 | > 30 | 29.8167 | 31.091 | 30.26 | 32.15 | 32.321 | > 30 | [5 empty] | #N/A
| 1 | 3 | 31.6 | 30.1032 | 29.81 | 30.71 | 31.32 | 31.51 | 30.2 | 32.3 | > 30 | 29.8453 | 30.493 | 30.74 | 32.18 | 32.559 | > 30 | [5 empty] | 16
| 1 | 4 | 32 | 30.1201 | 28.49 | 31.41 | 32.11 | 31.18 | 28.19 | 30.9 | > 30 | 29.892 | 30.988 | 30.23 | 32.14 | 32.235 | > 30 | [5 empty] | 31.177
| 1 | 5 | 31.4 | 30.624 | 29.79 | 31.51 | 31.51 | 32.3 | 28.84 | 31.51 | > 30 | 29.8209 | 31.9 | 31.54 | 32.18 | 32.027 | > 30 | [5 empty] | 111
| 1 | 6 | 30.9 | 29.9541 | 28.09 | 31.32 | 31.51 | 32.49 | 28.84 | 31.6 | > 30 | 29.8961 | 30.89 | 30.88 | 32.14 | 32.362 | > 30
| Mean |  | 31.3167 | 30.2526 | 28.89 | 31.2583 | 31.5767 | 31.8817 | 29.7233 | 31.6467 |  | 29.8479 | 31.056 | 30.82 | 32.1683 | 32.2688
| Median |  | 31.25 | 30.1582 | 28.58 | 31.365 | 31.51 | 31.905 | 29.52 | 31.645 |  | 29.8331 | 30.981 | 30.81 | 32.165 | 32.278
| Std Dev. |  | 0.43551 | 0.26088 | 0.72851 | 0.31064 | 0.2784 | 0.49053 | 1.359 | 0.46077 |  | 0.037338 | 0.46248 | 0.52775 | 0.031252 | 0.19014
| Rel.Std.Dev. |  | 1.391% | 0.8623% | 2.522% | 0.9938% | 0.8817% | 1.539% | 4.572% | 1.456% |  | 0.1251% | 1.489% | 1.712% | 0.09715% | 0.5892%
| PDM3 |  | 0.4481% | -2.965% | -7.335% | 0.261% | 1.282% | 2.26% | -4.663% | 1.507% |  | -4.263% | -0.388% | -1.145% | 3.18% | 3.502%
| Z-Score (Absolute) |  | 0.06 | 1.09 | 2.56 | 0 | 0.35 | 0.67 | 1.66 | 0.42 | NA | 1.53 | 0.22 | 0.47 | 0.99 | 1.09 | NA
| Analytical results for Sm in OREAS 906b (Certified Value 7.13 ppm) | [21 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 7.5 | 7.2 | 8.1 | 6.3 | 6.57 | 6.9 | 8.6 | 9 | [12 empty] | 1
| 1 | 2 | 7 | 7.1 | 8.2 | 6.3 | 6.73 | 7.3 | 5.8 | 9 | [12 empty] | 18
| 1 | 3 | 7.4 | 6.4 | 7.8 | 6.1 | 6.6 | 7.5 | 8.3 | 8 | [12 empty] | 16
| 1 | 4 | 7.5 | 6.9 | 7.8 | 6.4 | 6.72 | 7 | 6.7 | 8 | [12 empty] | 7.1319
| 1 | 5 | 7 | 7.6 | 8 | 6.6 | 6.62 | 10.3 | 7.7 | 8 | [12 empty] | 112
| 1 | 6 | 7 | 7.4 | 8.1 | 6 | 6.74 | 6.6 | 8.4 | 8
| Mean |  | 7.2333 | 7.1 | 8 | 6.2833 | 6.6633 | 7.6 | 7.5833 | 8.3333
| Median |  | 7.2 | 7.15 | 8.05 | 6.3 | 6.67 | 7.15 | 8 | 8
| Std Dev. |  | 0.2582 | 0.41952 | 0.16733 | 0.2137 | 0.075011 | 1.3594 | 1.1125 | 0.5164
| Rel.Std.Dev. |  | 3.57% | 5.909% | 2.092% | 3.401% | 1.126% | 17.89% | 14.67% | 6.197%
| PDM3 |  | 1.422% | -0.4474% | 12.17% | -11.9% | -6.57% | 6.563% | 6.33% | 16.85%
| Z-Score (Absolute) |  | 0 | 0.25 | 1.41 | 1.75 | 1.05 | 0.67 | 0.64 | NA
Results from laboratory 2.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 906b (Indicative Value 8.66 ppm) | [21 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.20 | 2.24 | 2.25 | 2.26 | 2.27 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | [9 empty] | 2
2
| 1 | 1 | 9 | < 100 | 21 | < 5 | 9 | 39 | 8.8 | < 50 | 8.6 | NA | 60 | [9 empty] | 1
| 1 | 2 | 9 | < 100 | 24 | < 5 | 9 | 40 | 8.1 | < 50 | 8.2 | NA | < 50 | [9 empty] | 10
| 1 | 3 | 9 | < 100 | < 5 | < 5 | 9 | 37 | 7.6 | < 50 | 8.5 | NA | < 50 | [9 empty] | 16
| 1 | 4 | 9 | < 100 | < 5 | < 5 | 8 | 39 | 8.6 | < 50 | 8.6 | NA | < 50 | [9 empty] | 8.6608
| 1 | 5 | 9 | < 100 | < 5 | < 5 | 8 | 38 | 9.2 | < 50 | 8.4 | NA | < 50 | [9 empty] | 16
| 1 | 6 | 9 | < 100 | < 5 | < 5 | 9 | 37 | 8.8 | < 50 | 9.4 | NA | 60
| Mean |  | 9 |  | 22.5 |  | 8.6667 | 38.3333 | 8.5167 |  | 8.6167 |  | 60
| Median |  | 9 |  | 22.5 |  | 9 | 38.5 | 8.7 |  | 8.55 |  | 60
| Std Dev. |  | 0 |  | 2.1213 |  | 0.5164 | 1.2111 | 0.57417 |  | 0.41191 |  | 0
| Rel.Std.Dev. |  | 0% |  | 9.428% |  | 5.958% | 3.159% | 6.742% |  | 4.78% |  | 0%
| PDM3 |  | 3.916% |  | 159.8% |  | 0.06735% | 342.6% | -1.665% |  | -0.51% |  | 592.8%
| Z-Score (Absolute) |  | 0.02 | 7.72 | 0.02 | 1.24 | 0.08 | 5.52 | 0.11 | 3 | 0.09 | NA | 5.2
| Analytical results for Sr in OREAS 906b (Certified Value 164 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 176 | 170 | 168.6659 | 170 | 11 | 163 | 179.6 | 151.73 | 57.4 | 151 | 165 | 134 | 169 | [7 empty] | 1
| 1 | 2 | 176 | 170 | 168.7721 | 170 | 11.2 | 166 | 178.2 | 153.91 | 59.6 | 155 | 165 | 136 | 157 | [7 empty] | 25
| 1 | 3 | 174 | 170 | 159.8834 | 170 | 11.1 | 159 | 171.9 | 153.59 | 58 | 151 | 164 | 134 | 165 | [7 empty] | 16
| 1 | 4 | 176 | 170 | 164.0663 | 170 | 11 | 174 | 174.2 | 159.49 | 58.3 | 155 | 165 | 137 | 170 | [7 empty] | 163.6181
| 1 | 5 | 177 | 180 | 164.4381 | 170 | 11 | 163 | 173.1 | 158.16 | 58.4 | 160 | 167 | 135 | 167 | [7 empty] | 113
| 1 | 6 | 174 | 170 | 159.5499 | 170 | 11.1 | 171 | 169.7 | 154.84 | 59.8 | 156 | 163 | 135 | 169
| Mean |  | 175.5 | 171.6667 | 164.2293 | 170 | 11.0667 | 166 | 174.45 | 155.2867 | 58.5833 | 154.6667 | 164.8333 | 135.1667 | 166.1667
| Median |  | 176 | 170 | 164.2522 | 170 | 11.05 | 164.5 | 173.65 | 154.375 | 58.35 | 155 | 165 | 135 | 168
| Std Dev. |  | 1.2247 | 4.0825 | 4.0293 | 0 | 0.08165 | 5.5857 | 3.7825 | 2.9512 | 0.9347 | 3.3862 | 1.3292 | 1.169 | 4.8339
| Rel.Std.Dev. |  | 0.6979% | 2.378% | 2.453% | 0% | 0.7378% | 3.365% | 2.168% | 1.9% | 1.596% | 2.189% | 0.8064% | 0.8649% | 2.909%
| PDM3 |  | 7.262% | 4.919% | 0.3735% | 3.9% | -93.24% | 1.456% | 6.62% | -5.092% | -64.2% | -5.471% | 0.7427% | -17.39% | 1.558%
| Z-Score (Absolute) |  | 0.75 | 0.48 | 0.04 | 0.36 | 10.86 | 0.08 | 0.68 | 0.67 | 7.5 | 0.72 | 0 | 2.1 | 0.09
| Analytical results for Ta in OREAS 906b (Certified Value 1.55 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [10 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [10 empty] | 2
2
| 1 | 1 | 1.6 | 1.1049 | < 2 | 2 | 1.4 | 10.8 | 1.21 | 1.69 | 2.1 | NA | [10 empty] | 1
| 1 | 2 | 1.4 | 1.0002 | < 2 | 2.1 | 1.4 | 11.4 | 1.25 | 1.64 | 2.2 | NA | [10 empty] | 36
| 1 | 3 | 1.5 | 0.93308 | < 2 | 2 | 1.5 | 11.6 | 1.24 | 1.42 | 2.6 | NA | [10 empty] | 16
| 1 | 4 | 1.7 | 1.1769 | < 2 | 1.9 | 1.4 | 12.2 | 1.21 | 1.33 | 1.7 | NA | [10 empty] | 1.5537
| 1 | 5 | 1.7 | 0.93285 | < 2 | 1.9 | 1.4 | 12.6 | 1.22 | 1.39 | 2.2 | NA | [10 empty] | 114
| 1 | 6 | 1.4 | 1.0895 | < 2 | 1.9 | 1.4 | 11.9 | 1.22 | 1.7 | 2.1 | NA
| Mean |  | 1.55 | 1.0396 |  | 1.9667 | 1.4167 | 11.75 | 1.225 | 1.5283 | 2.15
| Median |  | 1.55 | 1.0449 |  | 1.95 | 1.4 | 11.75 | 1.22 | 1.53 | 2.15
| Std Dev. |  | 0.13784 | 0.099895 |  | 0.08165 | 0.040825 | 0.63166 | 0.016432 | 0.1663 | 0.2881
| Rel.Std.Dev. |  | 8.893% | 9.609% |  | 4.152% | 2.882% | 5.376% | 1.341% | 10.88% | 13.4%
| PDM3 |  | -0.2412% | -33.09% |  | 26.58% | -8.823% | 656.2% | -21.16% | -1.636% | 38.38%
| Z-Score (Absolute) |  | 0.03 | 0.75 | 0.81 | 0.67 | 0.17 | 15.72 | 0.47 | 0 | 0.96 | NA
| Analytical results for Tb in OREAS 906b (Certified Value 0.79 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 0.8 | 0.9 | 0.8 | 0.83 | 0.68 | 0.82 | 0.8 | NA | [12 empty] | 1
| 1 | 2 | 0.7 | 0.9 | 0.84 | 0.81 | 0.74 | 0.8 | 0.6 | NA | [12 empty] | 37
| 1 | 3 | 0.8 | 0.8 | 0.83 | 0.77 | 0.72 | 0.75 | 0.9 | NA | [12 empty] | 16
| 1 | 4 | 0.8 | 0.9 | 0.8 | 0.76 | 0.73 | 0.73 | 0.6 | NA | [12 empty] | 0.78633
| 1 | 5 | 0.7 | 0.9 | 0.82 | 0.76 | 0.73 | 0.87 | 0.9 | NA | [12 empty] | 115
| 1 | 6 | 0.8 | 0.8 | 0.82 | 0.78 | 0.71 | 0.7 | 0.9 | NA
| Mean |  | 0.76667 | 0.86667 | 0.81833 | 0.785 | 0.71833 | 0.77833 | 0.78333
| Median |  | 0.8 | 0.9 | 0.82 | 0.775 | 0.725 | 0.775 | 0.85
| Std Dev. |  | 0.05164 | 0.05164 | 0.016021 | 0.02881 | 0.02137 | 0.063061 | 0.1472
| Rel.Std.Dev. |  | 6.736% | 5.958% | 1.958% | 3.67% | 2.975% | 8.102% | 18.79%
| PDM3 |  | -2.501% | 10.22% | 4.07% | -0.1696% | -8.648% | -1.017% | -0.3815%
| Z-Score (Absolute) |  | 0.67 | 3.37 | 1.42 | 0.07 | 2.63 | 0.2 | 0 | NA
| Analytical results for Te in OREAS 906b (Indicative Value 2.11 ppm) | [21 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 2.03 | 2.12 | 2.20 | 2.25 | [16 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES/MS | [16 empty] | 2
2
| 1 | 1 | < 1 | 0.3 | 0.3 | 7 | [16 empty] | 1
| 1 | 2 | < 1 | 0.4 | 0.2 | < 6 | [16 empty] | 11
| 1 | 3 | < 1 | 0.1 | 0.2 | 7 | [16 empty] | 16
| 1 | 4 | < 1 | 0.3 | 0.2 | < 6 | [16 empty] | 2.1111
| 1 | 5 | < 1 | 0.2 | 0.3 | 8 | [16 empty] | 17
| 1 | 6 | < 1 | 0.2 | 0.3 | 7
| Mean |  |  | 0.25 | 0.25 | 7.25
| Median |  |  | 0.25 | 0.25 | 7
| Std Dev. |  |  | 0.10488 | 0.054772 | 0.5
| Rel.Std.Dev. |  |  | 41.95% | 21.91% | 6.897%
| PDM3 |  |  | -88.16% | -88.16% | 243.4%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 29.44
| Analytical results for Th in OREAS 906b (Certified Value 14.7 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [11 empty] | 1
2
| 1 | 1 | 14.4 | 15 | 15.0796 | 15.6 | 14.8 | 11 | 10.54 | 14.1 | 15.3 | [11 empty] | 1
| 1 | 2 | 14.8 | 15 | 14.1535 | 15.2 | 15 | 11.9 | 10.95 | 15.1 | 13.5 | [11 empty] | 39
| 1 | 3 | 14.7 | 15.3 | 13.8147 | 14.4 | 14.7 | 11.4 | 10.42 | 15.3 | 14.8 | [11 empty] | 16
| 1 | 4 | 14.4 | 15.3 | 15.182 | 14.9 | 14.7 | 11.2 | 10.37 | 14.3 | 14.4 | [11 empty] | 14.7302
| 1 | 5 | 14.7 | 14.9 | 14.3281 | 15.4 | 14.8 | 10.2 | 10.54 | 14.4 | 14.7 | [11 empty] | 116
| 1 | 6 | 14.8 | 14.8 | 13.4109 | 14.9 | 14.8 | 10.5 | 10.52 | 13.4 | 14.8
| Mean |  | 14.6333 | 15.05 | 14.3281 | 15.0667 | 14.8 | 11.0333 | 10.5567 | 14.4333 | 14.5833
| Median |  | 14.7 | 15 | 14.2408 | 15.05 | 14.8 | 11.1 | 10.53 | 14.35 | 14.75
| Std Dev. |  | 0.18619 | 0.20736 | 0.69703 | 0.4274 | 0.10954 | 0.61536 | 0.205 | 0.69186 | 0.6047
| Rel.Std.Dev. |  | 1.272% | 1.378% | 4.865% | 2.837% | 0.7402% | 5.577% | 1.942% | 4.793% | 4.147%
| PDM3 |  | -0.6577% | 2.171% | -2.73% | 2.284% | 0.4738% | -25.1% | -28.33% | -2.015% | -0.9971%
| Z-Score (Absolute) |  | 0.13 | 1.23 | 0.67 | 1.28 | 0.57 | 9.38 | 10.64 | 0.4 | 0
| Analytical results for Ti in OREAS 906b (Certified Value 0.115 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 3
3
| 1 | 1 | 0.11 | 0.12 | 0.11125 | 0.12 | 0.02 | 0.156 | 0.114 | 0.114 | 0.114 | 0.114 | 0.12 | < 0.01 | 0.111 | 0.11 | 0.11 | 0.115 | 0.11 | 0.11 |  |  | 1
| 1 | 2 | 0.12 | 0.12 | 0.11181 | 0.12 | 0.02 | 0.132 | 0.114 | 0.12 | 0.12 | 0.114 | 0.12 | < 0.01 | 0.117 | 0.11 | 0.11 | 0.118 | 0.11 | 0.11 |  |  | 40
| 1 | 3 | 0.12 | 0.12 | 0.099006 | 0.13 | 0.02 | 0.114 | 0.108 | 0.162 | 0.114 | 0.114 | 0.12 | < 0.01 | 0.111 | 0.11 | 0.12 | 0.115 | 0.12 | 0.11 |  |  | 16
| 1 | 4 | 0.13 | 0.12 | 0.10797 | 0.12 | 0.02 | 0.12 | 0.108 | 0.114 | 0.108 | 0.108 | 0.12 | < 0.01 | 0.114 | 0.11 | 0.12 | 0.115 | 0.12 | 0.11 |  |  | 0.11488
| 1 | 5 | 0.12 | 0.12 | 0.11362 | 0.12 | 0.02 | 0.12 | 0.108 | 0.168 | 0.114 | 0.108 | 0.12 | < 0.01 | 0.113 | 0.11 | 0.11 | 0.121 | 0.11 | 0.11 |  |  | 117
| 1 | 6 | 0.12 | 0.13 | 0.10021 | 0.12 | 0.02 | 0.12 | 0.114 | 0.114 | 0.108 | 0.114 | 0.12 | < 0.01 | 0.116 | 0.11 | 0.11 | 0.114 | 0.11 | 0.11
| Mean |  | 0.12 | 0.12167 | 0.10731 | 0.12167 | 0.02 | 0.127 | 0.111 | 0.132 | 0.113 | 0.112 | 0.12 |  | 0.11367 | 0.11 | 0.11333 | 0.11633 | 0.11333 | 0.11
| Median |  | 0.12 | 0.12 | 0.10961 | 0.12 | 0.02 | 0.12 | 0.111 | 0.117 | 0.114 | 0.114 | 0.12 |  | 0.1135 | 0.11 | 0.11 | 0.115 | 0.11 | 0.11
| Std Dev. |  | 0.0063246 | 0.0040825 | 0.0062504 | 0.0040825 | 0 | 0.015375 | 0.0032863 | 0.025737 | 0.0045166 | 0.0030984 | 0 |  | 0.0025033 | 0 | 0.005164 | 0.0026583 | 0.005164 | 0
| Rel.Std.Dev. |  | 5.27% | 3.355% | 5.825% | 3.355% | 0% | 12.11% | 2.961% | 19.5% | 3.997% | 2.766% | 0% |  | 2.202% | 0% | 4.556% | 2.285% | 4.556% | 0%
| PDM3 |  | 4.458% | 5.909% | -6.588% | 5.909% | -82.59% | 10.55% | -3.377% | 14.9% | -1.636% | -2.506% | 4.458% |  | -1.055% | -4.247% | -1.345% | 1.266% | -1.345% | -4.247%
| Z-Score (Absolute) |  | 0.96 | 1.2 | 0.87 | 1.2 | 13.45 | 1.95 | 0.35 | 2.67 | 0.06 | 0.21 | 0.96 | 15.62 | 0.05 | 0.48 | 0 | 0.43 | 0 | 0.48
| Analytical results for Tl in OREAS 906b (Certified Value 0.74 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.20 | 2.25 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES/MS | [11 empty] | 2
3
| 1 | 1 | 0.7 | 0.8 | 0.761 | < 2 | < 1 | 0.7 | 0.6 | 0.7 | 0.7 | [11 empty] | 1
| 1 | 2 | 0.8 | 0.8 | 0.73952 | < 2 | < 1 | 0.7 | 0.6 | 0.8 | 0.8 | [11 empty] | 41
| 1 | 3 | 0.8 | 0.8 | 0.69456 | < 2 | < 1 | 0.7 | 0.7 | 0.8 | 0.7 | [11 empty] | 16
| 1 | 4 | 0.7 | 0.8 | 0.76153 | < 2 | < 1 | 0.7 | 0.7 | 0.8 | 0.7 | [11 empty] | 0.73883
| 1 | 5 | 0.8 | 0.8 | 0.69877 | < 2 | < 1 | 0.7 | 0.7 | 0.7 | 0.8 | [11 empty] | 118
| 1 | 6 | 0.8 | 0.8 | 0.67529 | < 2 | < 1 | 0.7 | 0.7 | 0.7 | 0.9
| Mean |  | 0.76667 | 0.8 | 0.72178 |  |  | 0.7 | 0.66667 | 0.75 | 0.76667
| Median |  | 0.8 | 0.8 | 0.71914 |  |  | 0.7 | 0.7 | 0.75 | 0.75
| Std Dev. |  | 0.05164 | 1.2162e-16 | 0.037051 |  |  | 1.2162e-16 | 0.05164 | 0.054772 | 0.08165
| Rel.Std.Dev. |  | 6.736% | 1.52e-14% | 5.133% |  |  | 1.737e-14% | 7.746% | 7.303% | 10.65%
| PDM3 |  | 3.768% | 8.28% | -2.307% |  |  | -5.255% | -9.767% | 1.512% | 3.768%
| Z-Score (Absolute) |  | 0.22 | 0.67 | 0.38 | 3.37 | 3.37 | 0.67 | 1.12 | 0 | 0.22
| Analytical results for Tm in OREAS 906b (Certified Value 0.12 ppm) | [21 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | 0.1 | < 1 | 0.13 | 0.09 | 0.11 | 0.13 | 0.1 | NA | [12 empty] | 1
| 1 | 2 | 0.1 | < 1 | 0.14 | 0.11 | 0.12 | 0.13 | 0.1 | NA | [12 empty] | 19
| 1 | 3 | 0.1 | < 1 | 0.15 | 0.1 | 0.12 | 0.13 | 0.2 | NA | [12 empty] | 16
| 1 | 4 | 0.1 | < 1 | 0.12 | 0.12 | 0.11 | 0.12 | 0.1 | NA | [12 empty] | 0.1225
| 1 | 5 | 0.1 | < 1 | 0.12 | 0.11 | 0.11 | 0.14 | 0.2 | NA | [12 empty] | 119
| 1 | 6 | 0.2 | < 1 | 0.13 | 0.12 | 0.12 | 0.13 | 0.1 | NA
| Mean |  | 0.11667 |  | 0.13167 | 0.10833 | 0.115 | 0.13 | 0.13333
| Median |  | 0.1 |  | 0.13 | 0.11 | 0.115 | 0.13 | 0.1
| Std Dev. |  | 0.040825 |  | 0.01169 | 0.01169 | 0.0054772 | 0.0063246 | 0.05164
| Rel.Std.Dev. |  | 34.99% |  | 8.879% | 10.79% | 4.763% | 4.865% | 38.73%
| PDM3 |  | -4.762% |  | 7.483% | -11.56% | -6.122% | 6.122% | 8.844%
| Z-Score (Absolute) |  | 0.67 | 18.71 | 0.08 | 1.1 | 0.76 | 0 | 0.17 | NA
| Analytical results for U in OREAS 906b (Certified Value 5.22 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | [11 empty] | 2
3
| 1 | 1 | 5.2 | 5.2 | 5.4335 | 5.3 | 5.47 | 4.71 | 2.3 | 5.04 | 5.8 | [11 empty] | 1
| 1 | 2 | 5.2 | 5.1 | 5.2099 | 5.2 | 5.52 | 4.8 | 2.4 | 5.23 | 4.8 | [11 empty] | 43
| 1 | 3 | 5.3 | 5.3 | 4.6504 | 5.1 | 5.46 | 4.58 | 2.4 | 4.94 | 5.3 | [11 empty] | 16
| 1 | 4 | 5.2 | 5.3 | 5.1882 | 5.1 | 5.46 | 4.69 | 2.3 | 4.97 | 5.1 | [11 empty] | 5.2216
| 1 | 5 | 5.4 | 5.1 | 5.3785 | 5.2 | 5.49 | 4.62 | 2.4 | 5.07 | 5.4 | [11 empty] | 120
| 1 | 6 | 5.3 | 5.2 | 4.5277 | 5.2 | 5.51 | 4.53 | 2.3 | 5.16 | 5.3
| Mean |  | 5.2667 | 5.2 | 5.0647 | 5.1833 | 5.485 | 4.655 | 2.35 | 5.0683 | 5.2833
| Median |  | 5.25 | 5.2 | 5.199 | 5.2 | 5.48 | 4.655 | 2.35 | 5.055 | 5.3
| Std Dev. |  | 0.08165 | 0.089443 | 0.38234 | 0.075277 | 0.025884 | 0.097724 | 0.054772 | 0.11089 | 0.33116
| Rel.Std.Dev. |  | 1.55% | 1.72% | 7.549% | 1.452% | 0.4719% | 2.099% | 2.331% | 2.188% | 6.268%
| PDM3 |  | 0.8626% | -0.4141% | -3.005% | -0.7333% | 5.044% | -10.85% | -54.99% | -2.936% | 1.182%
| Z-Score (Absolute) |  | 0.49 | 0.1 | 0.7 | 0 | 1.77 | 3.1 | 16.61 | 0.67 | 0.59
| Analytical results for V in OREAS 906b (Indicative Value 5.63 ppm) | [21 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.20 | 2.22 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 20 | < 20 | < 50 | 5.8 | < 5 | 7 | 6 | 5 | < 10 | 11 | [10 empty] | 1
| 1 | 2 | < 20 | < 20 | < 50 | 6 | < 5 | 5 | 5 | 5 | < 10 | 8 | [10 empty] | 12
| 1 | 3 | < 20 | < 20 | < 50 | 6 | < 5 | 5 | 5 | 5 | < 10 | 11 | [10 empty] | 16
| 1 | 4 | < 20 | < 20 | < 50 | 5.9 | < 5 | 7 | 5 | 5 | < 10 | 11 | [10 empty] | 5.6292
| 1 | 5 | < 20 | < 20 | < 50 | 6 | < 5 | 7 | 6 | 5 | < 10 | 11 | [10 empty] | 18
| 1 | 6 | < 20 | < 20 | < 50 | 5.8 | < 5 | 9 | < 5 | 5 | < 10 | 13
| Mean |  |  |  |  | 5.9167 |  | 6.6667 | 5.4 | 5 |  | 10.8333
| Median |  |  |  |  | 5.95 |  | 7 | 5 | 5 |  | 11
| Std Dev. |  |  |  |  | 0.098319 |  | 1.5055 | 0.54772 | 0 |  | 1.6021
| Rel.Std.Dev. |  |  |  |  | 1.662% |  | 22.58% | 10.14% | 0% |  | 14.79%
| PDM3 |  |  |  |  | 5.107% |  | 18.43% | -4.071% | -11.18% |  | 92.45%
| Z-Score (Absolute) |  | 0.98 | 0.98 | 4.96 | 0.1 | 1.01 | 0.1 | 0.36 | 0.34 | 0.34 | 1.2
| Analytical results for W in OREAS 906b (Indicative Value 2.79 ppm) | [21 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | 5 | 1.6986 | < 10 | < 5 | 3 | 3 | 1.5 | 3 | < 50 | 3.1 | [10 empty] | 1
| 1 | 2 | 5 | 1.7308 | < 10 | < 5 | 3 | 3 | 1.5 | 3 | < 50 | 1.8 | [10 empty] | 13
| 1 | 3 | 5 | 1.7598 | < 10 | < 5 | 3 | 4 | 1.5 | 2 | < 50 | 6.9 | [10 empty] | 16
| 1 | 4 | 4 | 1.7115 | < 10 | < 5 | 3 | 3 | 1.4 | 3 | < 50 | 2.4 | [10 empty] | 2.7944
| 1 | 5 | 4 | 1.7641 | < 10 | < 5 | 3 | 3 | 1.6 | 2 | < 50 | 2.8 | [10 empty] | 19
| 1 | 6 | 4 | 1.7618 | < 10 | < 5 | 3 | 4 | 1.4 | 3 | < 50 | 4.1
| Mean |  | 4.5 | 1.7378 |  |  | 3 | 3.3333 | 1.4833 | 2.6667 |  | 3.5167
| Median |  | 4.5 | 1.7453 |  |  | 3 | 3 | 1.5 | 3 |  | 2.95
| Std Dev. |  | 0.54772 | 0.028405 |  |  | 0 | 0.5164 | 0.075277 | 0.5164 |  | 1.8258
| Rel.Std.Dev. |  | 12.17% | 1.635% |  |  | 0% | 15.49% | 5.075% | 19.36% |  | 51.92%
| PDM3 |  | 61.03% | -37.81% |  |  | 7.356% | 19.28% | -46.92% | -4.573% |  | 25.84%
| Z-Score (Absolute) |  | 0.9 | 0.96 | 1.24 | 0.45 | 0.11 | 0.11 | 1.14 | 0.34 | 14.72 | 0.24
| Analytical results for Y in OREAS 906b (Certified Value 16.4 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | [8 empty] | 1
2
| 1 | 1 | 16.2 | 15.9 | 17.0066 | 15.5 | 17.1 | 17 | 12.8 | 17.9 | 15 | 12 | 18.9 | 18 | [8 empty] | 1
| 1 | 2 | 16.1 | 15.9 | 17.1221 | 14.9 | 17.5 | 16.8 | 13.4 | 15.6 | 15 | 11 | 13.5 | 17 | [8 empty] | 46
| 1 | 3 | 16.5 | 15.9 | 14.7138 | 13.2 | 17 | 17.1 | 13.1 | 19.4 | 15 | 11 | 17.2 | 16 | [8 empty] | 16
| 1 | 4 | 16.3 | 16.1 | 16.283 | 14.5 | 17 | 16.5 | 13.1 | 16.5 | 15 | 11 | 13.5 | 17 | [8 empty] | 16.3803
| 1 | 5 | 16.4 | 16.1 | 16.6309 | 15.7 | 17 | 16.3 | 12.9 | 17.6 | 15 | 11 | 16.8 | 17 | [8 empty] | 121
| 1 | 6 | 16.1 | 16.3 | 14.7993 | 14.5 | 17.3 | 16 | 13.3 | 20.3 | 15 | 11 | 17.8 | 18
| Mean |  | 16.2667 | 16.0333 | 16.0926 | 14.7167 | 17.15 | 16.6167 | 13.1 | 17.8833 | 15 | 11.1667 | 16.2833 | 17.1667
| Median |  | 16.25 | 16 | 16.457 | 14.7 | 17.05 | 16.65 | 13.1 | 17.75 | 15 | 11 | 17 | 17
| Std Dev. |  | 0.1633 | 0.1633 | 1.0766 | 0.89536 | 0.20736 | 0.42622 | 0.22804 | 1.752 | 0 | 0.40825 | 2.2693 | 0.75277
| Rel.Std.Dev. |  | 1.004% | 1.018% | 6.69% | 6.084% | 1.209% | 2.565% | 1.741% | 9.797% | 0% | 3.656% | 13.94% | 4.385%
| PDM3 |  | -0.6939% | -2.118% | -1.756% | -10.16% | 4.699% | 1.443% | -20.03% | 9.176% | -8.427% | -31.83% | -0.5921% | 4.801%
| Z-Score (Absolute) |  | 0 | 0.45 | 0.34 | 2.99 | 1.7 | 0.67 | 6.1 | 3.11 | NA | NA | 0.03 | NA
Results from laboratories 2.22, 2.24 and 2.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 906b (Certified Value 0.68 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.14 | 2.19 | 2.20 | 2.25 | 2.26 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [12 empty] | 2
3
| 1 | 1 | 0.7 | 0.7 | 0.7 | 0.7 | 0.66 | 0.7 | 0.7 | < 2 | [12 empty] | 1
| 1 | 2 | 0.5 | 0.7 | 0.7 | 0.7 | 0.68 | 0.8 | 0.7 | < 2 | [12 empty] | 47
| 1 | 3 | 0.7 | 0.7 | 0.7 | 0.6 | 0.69 | 0.6 | 0.5 | < 2 | [12 empty] | 16
| 1 | 4 | 0.7 | 0.7 | 0.7 | 0.7 | 0.62 | 0.7 | 0.6 | < 2 | [12 empty] | 0.67944
| 1 | 5 | 0.7 | 0.7 | 0.7 | 0.8 | 0.67 | 0.7 | 0.7 | < 2 | [12 empty] | 122
| 1 | 6 | 0.5 | 0.7 | 0.7 | 0.7 | 0.64 | 0.6 | 0.4 | < 2
| Mean |  | 0.63333 | 0.7 | 0.7 | 0.7 | 0.66 | 0.68333 | 0.6
| Median |  | 0.7 | 0.7 | 0.7 | 0.7 | 0.665 | 0.7 | 0.65
| Std Dev. |  | 0.10328 | 1.2162e-16 | 1.2162e-16 | 0.063246 | 0.026077 | 0.075277 | 0.12649
| Rel.Std.Dev. |  | 16.31% | 1.737e-14% | 1.737e-14% | 9.035% | 3.951% | 11.02% | 21.08%
| PDM3 |  | -6.787% | 3.025% | 3.025% | 3.025% | -2.862% | 0.5724% | -11.69%
| Z-Score (Absolute) |  | 1.97 | 0.28 | 0.28 | 0.28 | 1.07 | 0.28 | 3.09 | 10.4
| Analytical results for Zn in OREAS 906b (Certified Value 153 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 0
0
| 1 | 1 | 158 | 140 | 148.3097 | 126 | 68.8 | 140 | 150 | 280 | 190 | 120 | 164 | 145 | 41 | 155 | 165 | 146 | 150 | 100 |  |  | 1
| 1 | 2 | 155 | 150 | 147.3008 | 139 | 67.3 | 140 | 150 | 730 | 210 | 130 | 152 | 146 | 35 | 158 | 167 | 155 | 150 | 100 |  |  | 28
| 1 | 3 | 151 | 150 | 142.2489 | 125 | 65 | 140 | 150 | 260 | 30 | 120 | 148 | 142 | 37 | 134 | 165 | 152 | 160 | 100 |  |  | 16
| 1 | 4 | 155 | 150 | 142.7121 | 133 | 66.6 | 140 | 150 | 330 | 190 | 120 | 128 | 149 | 39 | 136 | 167 | 156 | 140 | 100 |  |  | 153.2494
| 1 | 5 | 155 | 140 | 150.4064 | 140 | 70.2 | 140 | 150 | 230 | 200 | 130 | 140 | 148 | 37 | 135 | 164 | 150 | 150 | 100 |  |  | 123
| 1 | 6 | 154 | 150 | 146.4877 | 124 | 67.1 | 140 | 150 | 310 | 220 | 130 | 131 | 137 | 38 | 139 | 164 | 160 | 160 | 100
| Mean |  | 154.6667 | 146.6667 | 146.2443 | 131.1667 | 67.5 | 140 | 150 | 356.6667 | 173.3333 | 125 | 143.8333 | 144.5 | 37.8333 | 142.8333 | 165.3333 | 153.1667 | 151.6667 | 100
| Median |  | 155 | 150 | 146.8943 | 129.5 | 67.2 | 140 | 150 | 295 | 195 | 125 | 144 | 145.5 | 37.5 | 137.5 | 165 | 153.5 | 150 | 100
| Std Dev. |  | 2.2509 | 5.164 | 3.2001 | 7.1949 | 1.8022 | 0 | 0 | 186.2973 | 71.1805 | 5.4772 | 13.5708 | 4.4159 | 2.0412 | 10.7595 | 1.3663 | 4.916 | 7.5277 | 0
| Rel.Std.Dev. |  | 1.455% | 3.521% | 2.188% | 5.485% | 2.67% | 0% | 0% | 52.23% | 41.07% | 4.382% | 9.435% | 3.056% | 5.395% | 7.533% | 0.8264% | 3.21% | 4.963% | 0%
| PDM3 |  | 0.9248% | -4.295% | -4.571% | -14.41% | -55.95% | -8.646% | -2.12% | 132.7% | 13.11% | -18.43% | -6.144% | -5.709% | -75.31% | -6.797% | 7.885% | -0.05396% | -1.033% | -34.75%
| Z-Score (Absolute) |  | 0.53 | NA | 0.05 | 0.82 | 4.47 | NA | NA | 12.13 | 1.6 | NA | 0.09 | 0.05 | 6.17 | 0.15 | 1.15 | 0.45 | NA | NA
Results from laboratories 2.02, 2.07, 2.08, 2.11 and 2.25 were removed due to their 10 ppm reading resolution. Results from laboratory 2.27 were removed due to their 100 ppm reading resolution.
| Analytical results for Zr in OREAS 906b (Indicative Value 278 ppm) | [21 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.19 | 2.26 | [15 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | [15 empty] | 0
0
| 1 | 1 | 284 | 58.3 | 268 | 278 | 281 | [15 empty] | 1
| 1 | 2 | 283 | 61.2 | 273 | 284.3 | 284 | [15 empty] | 14
| 1 | 3 | 286 | 57.9 | 271 | 284 | 275 | [15 empty] | 16
| 1 | 4 | 281 | 60.9 | 270 | 285.5 | 274 | [15 empty] | 278.4583
| 1 | 5 | 283 | 59 | 266 | 278.3 | 275 | [15 empty] | 20
| 1 | 6 | 284 | 58.2 | 272 | 283.9 | 279
| Mean |  | 283.5 | 59.25 | 270 | 282.3333 | 278
| Median |  | 283.5 | 58.65 | 270.5 | 283.95 | 277
| Std Dev. |  | 1.6432 | 1.4433 | 2.6077 | 3.2916 | 4
| Rel.Std.Dev. |  | 0.5796% | 2.436% | 0.9658% | 1.166% | 1.439%
| PDM3 |  | 1.811% | -78.72% | -3.038% | 1.392% | -0.1646%
| Z-Score (Absolute) |  | 0.67 | 26.82 | 0.98 | 0.53 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C869: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C887: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C905: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C923: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C959: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C977: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C995: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1013: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1031: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1104: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1122: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 906b (Certified Value 0.672 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 3
| 1 | 1 | 0.73 | 0.59 | 0.71528 | 0.582 | 0.55 | 0.68 | 0.67 | 0.72 | 0.68 | 0.63 | 0.63 | 0.6 | 0.65 | 0.66963 | < 2 | 0.7 |  | < 0.5 | 0.66 | 0.88 | 0.65 | 0.65 | 0.62 | 0.72 | 0.9 | 0.671 | 1
| 1 | 2 | 0.7 | 0.58 | 0.69923 | 0.695 | 0.54 | 0.7 | 0.72 | 0.72 | 0.69 | 0.64 | 0.63 | 0.2 | 0.66 | 0.71276 | < 2 | 0.7 |  | < 0.5 | 0.65 | 0.82 | 0.69 | 0.64 | 0.68 | 0.721 | 0.6 | 0.653 | 9
| 1 | 3 | 0.7 | 0.62 | 0.73513 | 0.772 | 0.55 | 0.72 | 0.72 | 0.7 | 0.7 | 0.64 | 0.64 | 0.4 | 0.67 | 0.67447 | < 2 | 0.7 |  | < 0.5 | 0.67 | 0.88 | 0.72 | 0.64 | 0.69 | 0.738 | 0.6 | 0.683 | 16
| 1 | 4 | 0.72 | 0.59 | 0.71436 | 0.599 | 0.56 | 0.69 | 0.68 | 0.71 | 0.66 | 0.63 | 0.65 | 0.5 | 0.64 | 0.72319 | < 2 | 0.8 |  | < 0.5 | 0.64 | 0.88 | 0.67 | 0.62 | 0.66 | 0.711 | 0.7 | 0.669 | 0.67162
| 1 | 5 | 0.7 | 0.58 | 0.72912 | 0.659 | 0.54 | 0.7 | 0.7 | 0.79 | 0.64 | 0.61 | 0.64 | 0.3 | 0.65 | 0.66419 | < 2 | 0.7 |  | < 0.5 | 0.71 | 0.84 | 0.67 | 0.65 | 0.62 | 0.715 | 0.6 | 0.674 | 125
| 1 | 6 | 0.72 | 0.57 | 0.70896 | 0.824 | 0.55 | 0.68 | 0.67 | 0.69 | 0.71 | 0.6 | 0.64 | 0.5 | 0.64 | 0.71371 | < 2 | 0.7 |  | < 0.5 | 0.67 | 0.97 | 0.69 | 0.67 | 0.57 | 0.722 | 0.6 | 0.681
| Mean |  | 0.71167 | 0.58833 | 0.71701 | 0.6885 | 0.54833 | 0.695 | 0.69333 | 0.72167 | 0.68 | 0.625 | 0.63833 | 0.41667 | 0.65167 | 0.693 |  | 0.71667 |  |  | 0.66667 | 0.87833 | 0.68167 | 0.645 | 0.64 | 0.72117 | 0.66667 | 0.67183
| Median |  | 0.71 | 0.585 | 0.71482 | 0.677 | 0.55 | 0.695 | 0.69 | 0.715 | 0.685 | 0.63 | 0.64 | 0.45 | 0.65 | 0.69362 |  | 0.7 |  |  | 0.665 | 0.88 | 0.68 | 0.645 | 0.64 | 0.7205 | 0.6 | 0.6725
| Std Dev. |  | 0.013292 | 0.017224 | 0.013158 | 0.095496 | 0.0075277 | 0.015166 | 0.023381 | 0.035449 | 0.026077 | 0.016432 | 0.0075277 | 0.1472 | 0.01169 | 0.02627 |  | 0.040825 |  |  | 0.024221 | 0.051543 | 0.024014 | 0.016432 | 0.045166 | 0.0092394 | 0.12111 | 0.010741
| Rel.Std.Dev. |  | 1.868% | 2.928% | 1.835% | 13.87% | 1.373% | 2.182% | 3.372% | 4.912% | 3.835% | 2.629% | 1.179% | 35.33% | 1.794% | 3.791% |  | 5.696% |  |  | 3.633% | 5.868% | 3.523% | 2.548% | 7.057% | 1.281% | 18.17% | 1.599%
| PDM3 |  | 5.963% | -12.4% | 6.76% | 2.514% | -18.36% | 3.482% | 3.234% | 7.452% | 1.248% | -6.941% | -4.956% | -37.96% | -2.97% | 3.183% |  | 6.708% |  |  | -0.7368% | 30.78% | 1.497% | -3.963% | -4.707% | 7.378% | -0.7368% | 0.03247%
| Z-Score (Absolute) |  | 0.58 | 1.73 | 0.68 | 0.14 | 2.48 | 0.27 | 0.23 | 0.76 | 0.02 | 1.05 | 0.8 | NA | 0.55 | 0.23 | 5.98 | NA | NA | 8.07 | 0.27 | 3.7 | 0.02 | 0.67 | 0.76 | 0.76 | NA | 0.17
Results from laboratories 2.13, 2.17 and 2.26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 906b (Certified Value 0.575 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.571 | 0.727 | 0.56497 | 0.58 | 3.54 | 0.56 | 0.55 | 0.63 | 0.53 | 0.59 | 0.51 | 0.534 | 0.52 | 0.59854 | 0.575 | 0.54 | 0.59 | 1.08 | 0.546 | 0.61 | 0.78 | 0.61 | 0.66 | 0.59 |  |  | 1
| 1 | 2 | 0.5882 | 0.732 | 0.5736 | 0.56 | 3.52 | 0.57 | 0.55 | 0.61 | 0.55 | 0.55 | 0.52 | 0.5 | 0.51 | 0.58119 | 0.584 | 0.52 | 0.58 | 0.968 | 0.559 | 0.61 | 0.76 | 0.65 | 0.71 | 0.58 |  |  | #N/A
| 1 | 3 | 0.5885 | 0.748 | 0.60677 | 0.66 | 3.55 | 0.58 | 0.55 | 0.61 | 0.54 | 0.55 | 0.58 | 0.51 | 0.53 | 0.6033 | 0.579 | 0.52 | 0.56 | 0.979 | 0.576 | 0.6 | 0.74625 | 0.61 | 0.68 | 0.59 |  |  | 16
| 1 | 4 | 0.5925 | 0.759 | 0.57393 | 0.6 | 3.53 | 0.58 | 0.53 | 0.61 | 0.55 | 0.55 | 0.58 | 0.506 | 0.52 | 0.59685 | 0.585 | 0.53 | 0.57 | 1 | 0.571 | 0.61 | 0.792 | 0.62 | 0.69 | 0.59 |  |  | 0.57522
| 1 | 5 | 0.564 | 0.749 | 0.60368 | 0.56 | 3.53 | 0.59 | 0.55 | 0.62 | 0.53 | 0.54 | 0.55 | 0.52 | 0.51 | 0.5938 | 0.584 | 0.53 | 0.58 | 1.07 | 0.595 | 0.61 | 0.756 | 0.6 | 0.66 | 0.57 |  |  | 126
| 1 | 6 | 0.5835 | 0.739 | 0.57112 | 0.69 | 3.53 | 0.56 | 0.53 | 0.61 | 0.55 | 0.56 | 0.57 | 0.55 | 0.52 | 0.58707 | 0.575 | 0.53 | 0.54 | 1.04 | 0.563 | 0.64 | 0.77175 | 0.59 | 0.68 | 0.58
| Mean |  | 0.58128 | 0.74233 | 0.58234 | 0.60833 | 3.5333 | 0.57333 | 0.54333 | 0.615 | 0.54167 | 0.55667 | 0.55167 | 0.52 | 0.51833 | 0.59346 | 0.58033 | 0.52833 | 0.57 | 1.0228 | 0.56833 | 0.61333 | 0.76767 | 0.61333 | 0.68 | 0.58333
| Median |  | 0.58585 | 0.7435 | 0.57376 | 0.59 | 3.53 | 0.575 | 0.55 | 0.61 | 0.545 | 0.55 | 0.56 | 0.515 | 0.52 | 0.59533 | 0.5815 | 0.53 | 0.575 | 1.02 | 0.567 | 0.61 | 0.76588 | 0.61 | 0.68 | 0.585
| Std Dev. |  | 0.01127 | 0.011894 | 0.018037 | 0.054559 | 0.010328 | 0.012111 | 0.010328 | 0.0083666 | 0.0098319 | 0.017512 | 0.030605 | 0.018931 | 0.0075277 | 0.0080694 | 0.0046332 | 0.0075277 | 0.017889 | 0.047406 | 0.016681 | 0.013663 | 0.016819 | 0.020656 | 0.018974 | 0.008165
| Rel.Std.Dev. |  | 1.939% | 1.602% | 3.097% | 8.969% | 0.2923% | 2.112% | 1.901% | 1.36% | 1.815% | 3.146% | 5.548% | 3.641% | 1.452% | 1.36% | 0.7984% | 1.425% | 3.138% | 4.635% | 2.935% | 2.228% | 2.191% | 3.368% | 2.79% | 1.4%
| PDM3 |  | 1.054% | 29.05% | 1.238% | 5.756% | 514.3% | -0.3282% | -5.544% | 6.915% | -5.833% | -3.226% | -4.095% | -9.6% | -9.89% | 3.17% | 0.8888% | -8.151% | -0.9076% | 77.82% | -1.197% | 6.626% | 33.46% | 6.626% | 18.22% | 1.41%
| Z-Score (Absolute) |  | 0.01 | 3.43 | 0.01 | 0.57 | 63.14 | 0.18 | 0.82 | 0.71 | 0.86 | 0.54 | 0.64 | 1.32 | 1.36 | 0.25 | 0.03 | 1.14 | 0.25 | 9.43 | 0.29 | 0.67 | 3.98 | 0.67 | 2.1 | 0.03
| Analytical results for As in OREAS 906b (Certified Value 11.8 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 2
| 1 | 1 | 12 | 12 | 11.2152 | 12.8 | 7.5 | 11.4 | 11.8 | 11.8 | 11.5 | 11.3 | 11 | 14 | 11 | 12.0073 | 15 | 12.2 | 13.23 | 12 | 12.4 | 11.7 | 21.7 | 13 | 10.9 | 13.4 | 10.8 |  | 1
| 1 | 2 | 12 | 12 | 11.4573 | 12.8 | 7.1 | 11.7 | 12.1 | 11.7 | 11.8 | 11 | 11 | 13 | 12 | 12.2862 | 15 | 11.6 | 12.29 | 11 | 12.9 | 12 | 21.3 | 12 | 11.1 | 13.1 | 10.8 |  | 10
| 1 | 3 | 12 | 12 | 12.176 | 12 | 6.3 | 12.1 | 11.8 | 12 | 11.5 | 11.7 | 11 | 13 | 12 | 12.5112 | 13 | 11.4 | 11.88 | 12 | 13.1 | 11.6 | 21.3 | 13 | 11 | 12.6 | 10.9 |  | 16
| 1 | 4 | 12 | 12 | 11.6434 | 13 | 6.9 | 11.5 | 11.8 | 11.7 | 11.4 | 11 | 11 | 12 | 11 | 12.3181 | 18 | 11.7 | 11.71 | 12 | 12.1 | 11.8 | 22.2 | 13 | 10.6 | 12.8 | 11 |  | 11.8404
| 1 | 5 | 11 | 12 | 12.2507 | 12.8 | 6.3 | 11.6 | 11.8 | 12.1 | 11.3 | 11.4 | 10 | 12 | 12 | 12.4106 | 15 | 12 | 11.28 | 12 | 13.2 | 11.9 | 22.1 | 12 | 11.2 | 13.6 | 10.7 |  | 127
| 1 | 6 | 12 | 12 | 11.5284 | 12.2 | 6.9 | 10.9 | 11.8 | 11.7 | 11.7 | 11.2 | 11 | 15 | 12 | 11.9013 | 15 | 11.9 | 11.78 | 11 | 13 | 12.2 | 21.7 | 12 | 10.9 | 12.9 | 10.5
| Mean |  | 11.8333 | 12 | 11.7118 | 12.6 | 6.8333 | 11.5333 | 11.85 | 11.8333 | 11.5333 | 11.2667 | 10.8333 | 13.1667 | 11.6667 | 12.2391 | 15.1667 | 11.8 | 12.0283 | 11.6667 | 12.7833 | 11.8667 | 21.7167 | 12.5 | 10.95 | 13.0667 | 10.7833
| Median |  | 12 | 12 | 11.5859 | 12.8 | 6.9 | 11.55 | 11.8 | 11.75 | 11.5 | 11.25 | 11 | 13 | 12 | 12.3022 | 15 | 11.8 | 11.83 | 12 | 12.95 | 11.85 | 21.7 | 12.5 | 10.95 | 13 | 10.8
| Std Dev. |  | 0.40825 | 0 | 0.41366 | 0.4 | 0.46762 | 0.39328 | 0.12247 | 0.17512 | 0.18619 | 0.26583 | 0.40825 | 1.169 | 0.5164 | 0.23654 | 1.6021 | 0.28983 | 0.67194 | 0.5164 | 0.43551 | 0.21602 | 0.38166 | 0.54772 | 0.20736 | 0.37771 | 0.17224
| Rel.Std.Dev. |  | 3.45% | 0% | 3.532% | 3.175% | 6.843% | 3.41% | 1.034% | 1.48% | 1.614% | 2.359% | 3.768% | 8.879% | 4.426% | 1.933% | 10.56% | 2.456% | 5.586% | 4.426% | 3.407% | 1.82% | 1.757% | 4.382% | 1.894% | 2.891% | 1.597%
| PDM3 |  | -0.05948% | 1.348% | -1.086% | 6.416% | -42.29% | -2.593% | 0.08129% | -0.05948% | -2.593% | -4.845% | -8.505% | 11.2% | -1.467% | 3.368% | 28.09% | -0.341% | 1.587% | -1.467% | 7.964% | 0.222% | 83.41% | 5.571% | -7.52% | 10.36% | -8.927%
| Z-Score (Absolute) |  | NA | NA | 0.2 | 1.27 | 8.31 | 0.5 | 0.03 | 0 | 0.5 | 0.94 | NA | NA | NA | 0.67 | NA | 0.06 | 0.32 | NA | 1.58 | 0.06 | 16.42 | NA | 1.47 | 2.05 | 1.74
Results from laboratories 2.01, 2.02, 2.12, 2.13, 2.14, 2.16, 2.19 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 906b (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.22 | 2.23 | 2.25 | 2.27 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 30g | [13 empty] | 1
| 1 | 1 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 2 | < 10 | NA | 3 | 2 | [13 empty] | 1
| 1 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 2 | < 10 | NA | 2 | 2 | [13 empty] | 7
| 1 | 3 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 2 | < 10 | NA | 3 | 2 | [13 empty] | 16
| 1 | 4 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 2 | < 10 | NA | 3 | 2 | [13 empty] | < 10
| 1 | 5 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 2 | < 10 | NA | 3 | 2 | [13 empty] | 128
| 1 | 6 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 2 | < 10 | NA | 3 | 2
| Mean | [12 empty] | 2.8333 | 2
| Median | [12 empty] | 3 | 2
| Std Dev. | [12 empty] | 0.40825 | 0
| Rel.Std.Dev. | [12 empty] | 14.41% | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 906b (Certified Value 185 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 0
| 1 | 1 | 177 | 208 | 194.6359 | 176.3 | 187 | 180 | 180 | 200 | 150 | 190 | 168 | 158 | 164 | 183.1289 | 156 | 167 | 194 | 273 | 175 | 204 | 238 | 208 | 197 | 173.7 |  |  | 1
| 1 | 2 | 176 | 207 | 196.6132 | 177.4 | 189.1 | 180 | 180 | 200 | 150 | 170 | 171 | 160 | 168 | 178.5554 | 157 | 158 | 193 | 251 | 179 | 213 | 235 | 210 | 206 | 179.3 |  |  | 35
| 1 | 3 | 174 | 208 | 209.504 | 220.1 | 189 | 190 | 170 | 190 | 150 | 170 | 173 | 160 | 167 | 183.3606 | 155 | 157 | 189 | 250 | 184 | 208 | 238 | 202 | 208 | 181.3 |  |  | 16
| 1 | 4 | 181 | 213 | 198.7466 | 178.5 | 186.2 | 180 | 170 | 200 | 150 | 180 | 181 | 154 | 163 | 182.9979 | 159 | 164 | 188 | 261 | 178 | 204 | 241 | 172 | 204 | 173.6 |  |  | 184.8151
| 1 | 5 | 170 | 208 | 210.1085 | 173.8 | 186.3 | 190 | 180 | 200 | 150 | 170 | 176 | 160 | 165 | 184.0311 | 158 | 164 | 191 | 277 | 189 | 217 | 247 | 209 | 193 | 177.1 |  |  | 129
| 1 | 6 | 175 | 211 | 200.0216 | 229.6 | 192.6 | 180 | 170 | 200 | 150 | 180 | 178 | 160 | 164 | 182.0646 | 157 | 164 | 183 | 268 | 181 | 210 | 246 | 206 | 206 | 173.6
| Mean |  | 175.5 | 209.1667 | 201.605 | 192.6167 | 188.3667 | 183.3333 | 175 | 198.3333 | 150 | 176.6667 | 174.5 | 158.6667 | 165.1667 | 182.3564 | 157 | 162.3333 | 189.6667 | 263.3333 | 181 | 209.3333 | 240.8333 | 201.1667 | 202.3333 | 176.4333
| Median |  | 175.5 | 208 | 199.3841 | 177.95 | 188 | 180 | 175 | 200 | 150 | 175 | 174.5 | 160 | 164.5 | 183.0634 | 157 | 164 | 190 | 264.5 | 180 | 209 | 239.5 | 207 | 205 | 175.4
| Std Dev. |  | 3.6194 | 2.3166 | 6.6162 | 25.1961 | 2.4369 | 5.164 | 5.4772 | 4.0825 | 0 | 8.165 | 4.7645 | 2.4221 | 1.9408 | 1.9675 | 1.4142 | 3.9328 | 3.9833 | 11.2901 | 4.9396 | 5.1251 | 4.7924 | 14.5659 | 5.9554 | 3.3429
| Rel.Std.Dev. |  | 2.062% | 1.108% | 3.282% | 13.08% | 1.294% | 2.817% | 3.13% | 2.058% | 0% | 4.622% | 2.73% | 1.527% | 1.175% | 1.079% | 0.9008% | 2.423% | 2.1% | 4.287% | 2.729% | 2.448% | 1.99% | 7.241% | 2.943% | 1.895%
| PDM3 |  | -5.04% | 13.18% | 9.085% | 4.221% | 1.922% | -0.8018% | -5.311% | 7.314% | -18.84% | -4.409% | -5.581% | -14.15% | -10.63% | -1.33% | -15.05% | -12.16% | 2.625% | 42.48% | -2.064% | 13.27% | 30.31% | 8.848% | 9.479% | -4.535%
| Z-Score (Absolute) |  | 0.3 | 1.07 | 0.76 | 0.4 | 0.22 | 0.02 | 0.32 | 0.63 | 1.34 | 0.25 | 0.34 | 0.98 | 0.72 | 0.02 | 1.05 | 0.83 | 0.28 | 3.27 | 0.08 | 1.08 | 2.36 | 0.74 | 0.79 | 0.26
| Analytical results for Be in OREAS 906b (Certified Value 0.85 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.88 | 0.95 | 0.82406 | < 2 | 0.83 | 0.87 | 0.91 | 0.79 | 0.84 | 0.86 | < 0.5 | 0.8 | 0.8529 | 0.8 | 0.9 | 1.05 | 0.68 | 0.79 | 0.9 | 0.9 | [6 empty] | 1
| 1 | 2 | 0.86 | 0.94 | 0.76339 | < 2 | 0.84 | 0.86 | 0.88 | 0.81 | 0.82 | 0.86 | < 0.5 | 0.9 | 0.81193 | 0.8 | 0.7 | 1 | 0.68 | 0.78 | 0.9 | 0.8 | [6 empty] | 36
| 1 | 3 | 0.85 | 0.91 | 0.82409 | < 2 | 0.86 | 0.86 | 0.87 | 0.8 | 0.82 | 0.88 | < 0.5 | 0.8 | 0.81561 | 0.8 | 0.8 | 1.04 | 0.67 | 0.85 | 0.9 | 0.9 | [6 empty] | 16
| 1 | 4 | 0.87 | 0.93 | 0.82868 | < 2 | 0.86 | 0.82 | 0.88 | 0.82 | 0.82 | 0.9 | < 0.5 | 0.8 | 0.85532 | 0.8 | 0.8 | 1.02 | 0.66 | 0.74 | 0.9 | 0.8 | [6 empty] | 0.84692
| 1 | 5 | 0.84 | 0.91 | 0.77425 | < 2 | 0.86 | 0.87 | 0.9 | 0.8 | 0.82 | 0.85 | < 0.5 | 0.9 | 0.84778 | 0.8 | 0.8 | 1.05 | 0.6 | 0.82 | 0.9 | 0.8 | [6 empty] | 130
| 1 | 6 | 0.85 | 0.92 | 0.8393 | < 2 | 0.83 | 0.85 | 0.89 | 0.82 | 0.83 | 0.85 | < 0.5 | 0.8 | 0.81728 | 0.8 | 0.7 | 1.06 | 0.61 | 0.78 | 0.9 | 0.8
| Mean |  | 0.85833 | 0.92667 | 0.80896 |  | 0.84667 | 0.855 | 0.88833 | 0.80667 | 0.825 | 0.86667 |  | 0.83333 | 0.83347 | 0.8 | 0.78333 | 1.0367 | 0.65 | 0.79333 | 0.9 | 0.83333
| Median |  | 0.855 | 0.925 | 0.82408 |  | 0.85 | 0.86 | 0.885 | 0.805 | 0.82 | 0.86 |  | 0.8 | 0.83253 | 0.8 | 0.8 | 1.045 | 0.665 | 0.785 | 0.9 | 0.8
| Std Dev. |  | 0.01472 | 0.01633 | 0.031775 |  | 0.015055 | 0.018708 | 0.01472 | 0.012111 | 0.0083666 | 0.019664 |  | 0.05164 | 0.020515 | 1.2162e-16 | 0.075277 | 0.022509 | 0.035777 | 0.037771 | 0 | 0.05164
| Rel.Std.Dev. |  | 1.715% | 1.762% | 3.928% |  | 1.778% | 2.188% | 1.657% | 1.501% | 1.014% | 2.269% |  | 6.197% | 2.461% | 1.52e-14% | 9.61% | 2.171% | 5.504% | 4.761% | 0% | 6.197%
| PDM3 |  | 1.348% | 9.416% | -4.482% |  | -0.02967% | 0.9543% | 4.89% | -4.753% | -2.588% | 2.332% |  | -1.604% | -1.588% | -5.54% | -7.508% | 22.4% | -23.25% | -6.327% | 6.268% | -1.604%
| Z-Score (Absolute) |  | 0.2 | 1.35 | 0.64 | 2.58 | 0 | 0.14 | 0.7 | 0.67 | 0.37 | 0.34 | 10.06 | NA | 0.22 | NA | NA | 3.2 | 3.32 | 0.9 | NA | NA
Results from laboratories 2.14, 2.16, 2.17, 2.24 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 906b (Certified Value 9.91 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 9.96 | 9.91 | 9.2747 | 8.45 | < 5 | 9.16 | 9.29 | 11.1 | 9.98 | 9.28 | 10.23 | 10 | 9.53 | 9.3065 | 11 | 9.66 | 10 | 10.7 | 11.2 | 10.26 | 11.2 | 9.68 | 10 | 10.19 | 9.59 |  | 1
| 1 | 2 | 10.09 | 9.68 | 9.1645 | 9.4 | < 5 | 9.39 | 9.25 | 10.95 | 10.15 | 9.64 | 10.42 | < 2 | 9.36 | 9.7718 | 13 | 9.58 | 11 | 10.6 | 10.36 | 10.65 | 11.3 | 9.82 | 10.1 | 10.09 | 9.81 |  | 37
| 1 | 3 | 9.88 | 9.65 | 9.8827 | 10.05 | < 5 | 9.87 | 9.24 | 10.75 | 10.1 | 9.54 | 10.5 | 7 | 9.56 | 9.7517 | 11 | 9.18 | 12 | 10.8 | 10.26 | 10.25 | 12 | 9.68 | 10.2 | 9.76 | 9.77 |  | 16
| 1 | 4 | 10.16 | 10 | 9.5929 | 8.4 | < 5 | 9.55 | 9.47 | 10.95 | 10.15 | 9.34 | 10.5 | 7 | 9.33 | 9.5 | 12 | 9.82 | 11 | 10.6 | 11.54 | 10.32 | 11.8 | 9.68 | 9.98 | 10.11 | 9.49 |  | 9.9095
| 1 | 5 | 9.93 | 9.83 | 9.9679 | 8.89 | < 5 | 9.61 | 9.32 | 10.85 | 10 | 9.52 | 10.26 | 6 | 9.44 | 10.2203 | 12 | 9.51 | 11 | 10.7 | 10.61 | 9.97 | 12.9 | 9.86 | 10 | 10.07 | 9.89 |  | 131
| 1 | 6 | 9.99 | 9.81 | 9.732 | 9.61 | < 5 | 8.84 | 9.53 | 11 | 10.3 | 9.33 | 10.36 | 10 | 9.48 | 9.4905 | 11 | 9.58 | 10 | 10.4 | 10.96 | 10.23 | 11.7 | 9.64 | 10.1 | 10.11 | 9.65
| Mean |  | 10.0017 | 9.8133 | 9.6025 | 9.1333 |  | 9.4033 | 9.35 | 10.9333 | 10.1133 | 9.4417 | 10.3783 | 8 | 9.45 | 9.6734 | 11.6667 | 9.555 | 10.8333 | 10.6333 | 10.8217 | 10.28 | 11.8167 | 9.7267 | 10.0633 | 10.055 | 9.7
| Median |  | 9.975 | 9.82 | 9.6625 | 9.145 |  | 9.47 | 9.305 | 10.95 | 10.125 | 9.43 | 10.39 | 7 | 9.46 | 9.6258 | 11.5 | 9.58 | 11 | 10.65 | 10.785 | 10.255 | 11.75 | 9.68 | 10.05 | 10.1 | 9.71
| Std Dev. |  | 0.10458 | 0.13337 | 0.32503 | 0.66371 |  | 0.36297 | 0.12116 | 0.12111 | 0.1169 | 0.14428 | 0.11635 | 1.8708 | 0.091652 | 0.32031 | 0.8165 | 0.2122 | 0.75277 | 0.13663 | 0.50034 | 0.21836 | 0.61128 | 0.090037 | 0.085245 | 0.15017 | 0.14953
| Rel.Std.Dev. |  | 1.046% | 1.359% | 3.385% | 7.267% |  | 3.86% | 1.296% | 1.108% | 1.156% | 1.528% | 1.121% | 23.39% | 0.9699% | 3.311% | 6.999% | 2.221% | 6.949% | 1.285% | 4.623% | 2.124% | 5.173% | 0.9257% | 0.8471% | 1.493% | 1.542%
| PDM3 |  | 0.9305% | -0.9701% | -3.098% | -7.832% |  | -5.108% | -5.646% | 10.33% | 2.057% | -4.721% | 4.732% | -19.27% | -4.637% | -2.382% | 17.73% | -3.577% | 9.323% | 7.305% | 9.205% | 3.739% | 19.25% | -1.845% | 1.553% | 1.469% | -2.114%
| Z-Score (Absolute) |  | 0.47 | 0.09 | 0.34 | 1.28 | 14.6 | 0.74 | 0.84 | 2.34 | 0.69 | 0.66 | 1.22 | NA | 0.64 | 0.19 | NA | 0.43 | NA | 1.73 | 2.11 | 1.02 | 4.11 | 0.09 | 0.59 | 0.57 | 0.14
Results from laboratories 2.13, 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 906b (Certified Value 0.363 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 0.35 | 0.4 | 0.34951 | 0.33 | 0.57 | 0.37 | 0.36 | 0.39 | 0.37 | 0.37 | 0.357 | 0.389 | 0.37 | 0.35755 | 0.351 | 0.36 | 0.36 | 0.336 | 0.323 | 0.36 | 0.3112 | 0.33 | 0.37 | 0.3741 | 0.36 |  | 1
| 1 | 2 | 0.36 | 0.39 | 0.35352 | 0.35 | 0.56 | 0.38 | 0.36 | 0.38 | 0.38 | 0.37 | 0.365 | 0.39 | 0.36 | 0.35653 | 0.355 | 0.36 | 0.36 | 0.309 | 0.336 | 0.35 | 0.3149 | 0.34 | 0.39 | 0.3677 | 0.36 |  | #N/A
| 1 | 3 | 0.36 | 0.4 | 0.36422 | 0.39 | 0.57 | 0.38 | 0.36 | 0.38 | 0.38 | 0.37 | 0.362 | 0.4 | 0.37 | 0.3596 | 0.352 | 0.35 | 0.36 | 0.319 | 0.337 | 0.35 | 0.3156 | 0.33 | 0.38 | 0.3687 | 0.36 |  | 16
| 1 | 4 | 0.36 | 0.4 | 0.35795 | 0.35 | 0.56 | 0.38 | 0.35 | 0.38 | 0.38 | 0.36 | 0.362 | 0.399 | 0.37 | 0.35918 | 0.359 | 0.36 | 0.36 | 0.307 | 0.327 | 0.35 | 0.3167 | 0.33 | 0.37 | 0.3718 | 0.36 |  | 0.36288
| 1 | 5 | 0.36 | 0.4 | 0.36969 | 0.35 | 0.57 | 0.38 | 0.36 | 0.39 | 0.37 | 0.37 | 0.357 | 0.4 | 0.37 | 0.35906 | 0.355 | 0.36 | 0.36 | 0.37 | 0.342 | 0.35 | 0.3149 | 0.32 | 0.38 | 0.3838 | 0.35 |  | 132
| 1 | 6 | 0.36 | 0.4 | 0.36236 | 0.37 | 0.57 | 0.37 | 0.36 | 0.39 | 0.38 | 0.37 | 0.359 | 0.4 | 0.38 | 0.3599 | 0.358 | 0.36 | 0.36 | 0.319 | 0.338 | 0.35 | 0.3128 | 0.33 | 0.37 | 0.3641 | 0.36
| Mean |  | 0.35833 | 0.39833 | 0.35954 | 0.35667 | 0.56667 | 0.37667 | 0.35833 | 0.385 | 0.37667 | 0.36833 | 0.36033 | 0.39633 | 0.37 | 0.35864 | 0.355 | 0.35833 | 0.36 | 0.32667 | 0.33383 | 0.35167 | 0.31435 | 0.33 | 0.37667 | 0.3717 | 0.35833
| Median |  | 0.36 | 0.4 | 0.36016 | 0.35 | 0.57 | 0.38 | 0.36 | 0.385 | 0.38 | 0.37 | 0.3605 | 0.3995 | 0.37 | 0.35912 | 0.355 | 0.36 | 0.36 | 0.319 | 0.3365 | 0.35 | 0.3149 | 0.33 | 0.375 | 0.37025 | 0.36
| Std Dev. |  | 0.0040825 | 0.0040825 | 0.0073802 | 0.020656 | 0.005164 | 0.005164 | 0.0040825 | 0.0054772 | 0.005164 | 0.0040825 | 0.0032042 | 0.0053166 | 0.0063246 | 0.0013144 | 0.0031623 | 0.0040825 | 6.0809e-17 | 0.023585 | 0.0072503 | 0.0040825 | 0.0020007 | 0.0063246 | 0.008165 | 0.0068548 | 0.0040825
| Rel.Std.Dev. |  | 1.139% | 1.025% | 2.053% | 5.791% | 0.9113% | 1.371% | 1.139% | 1.423% | 1.371% | 1.108% | 0.8892% | 1.341% | 1.709% | 0.3665% | 0.8908% | 1.139% | 1.689e-14% | 7.22% | 2.172% | 1.161% | 0.6365% | 1.917% | 2.168% | 1.844% | 1.139%
| PDM3 |  | -1.252% | 9.771% | -0.9186% | -1.711% | 56.16% | 3.801% | -1.252% | 6.097% | 3.801% | 1.504% | -0.7005% | 9.22% | 1.963% | -1.168% | -2.17% | -1.252% | -0.7923% | -9.978% | -8.003% | -3.089% | -13.37% | -9.06% | 3.801% | 2.432% | -1.252%
| Z-Score (Absolute) |  | 0.08 | 2.5 | 0 | 0.19 | 13.35 | 1.1 | 0.08 | 1.64 | 1.1 | 0.57 | 0.05 | 2.37 | 0.67 | 0.06 | 0.29 | 0.08 | 0.03 | 2.12 | 1.66 | 0.51 | 2.91 | 1.9 | 1.1 | 0.78 | 0.08
| Analytical results for Cd in OREAS 906b (Certified Value 0.4 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 0.4 | 0.37 | 0.41834 | 0.35 | < 1 | 0.37 | 0.39 | 0.4 | 0.39 | 0.4 | 0.4 | < 0.5 | 0.41 | 0.39723 | < 1 | 0.45 | < 1 | 0.29 | 0.42 | 0.43 | 0.42 | 0.41 | 0.41 |  |  |  | 1
| 1 | 2 | 0.37 | 0.36 | 0.41142 | 0.41 | < 1 | 0.41 | 0.39 | 0.4 | 0.42 | 0.4 | 0.41 | < 0.5 | 0.41 | 0.37039 | < 1 | 0.45 | < 1 | 0.31 | 0.43 | 0.44 | 0.41 | 0.39 | 0.38 |  |  |  | 12
| 1 | 3 | 0.38 | 0.38 | 0.4508 | 0.5 | < 1 | 0.42 | 0.4 | 0.39 | 0.41 | 0.38 | 0.42 | < 0.5 | 0.4 | 0.37399 | < 1 | 0.43 | < 1 | 0.31 | 0.45 | 0.4 | 0.42 | 0.43 | 0.37 |  |  |  | 16
| 1 | 4 | 0.4 | 0.4 | 0.42697 | 0.35 | < 1 | 0.38 | 0.43 | 0.41 | 0.37 | 0.41 | 0.4 | < 0.5 | 0.39 | 0.35118 | < 1 | 0.49 | < 1 | 0.31 | 0.46 | 0.4 | 0.4 | 0.42 | 0.4 |  |  |  | 0.40096
| 1 | 5 | 0.4 | 0.37 | 0.45094 | 0.36 | < 1 | 0.39 | 0.4 | 0.42 | 0.4 | 0.37 | 0.4 | < 0.5 | 0.4 | 0.35879 | < 1 | 0.44 | < 1 | 0.3 | 0.41 | 0.4 | 0.44 | 0.4 | 0.41 |  |  |  | 133
| 1 | 6 | 0.42 | 0.39 | 0.4406 | 0.49 | < 1 | 0.4 | 0.42 | 0.43 | 0.43 | 0.39 | 0.4 | < 0.5 | 0.41 | 0.36614 | < 1 | 0.45 | < 1 | 0.31 | 0.41 | 0.4 | 0.42 | 0.42 | 0.4
| Mean |  | 0.395 | 0.37833 | 0.43318 | 0.41 |  | 0.395 | 0.405 | 0.40833 | 0.40333 | 0.39167 | 0.405 |  | 0.40333 | 0.36962 |  | 0.45167 |  | 0.305 | 0.43 | 0.41167 | 0.41833 | 0.41167 | 0.395
| Median |  | 0.4 | 0.375 | 0.43379 | 0.385 |  | 0.395 | 0.4 | 0.405 | 0.405 | 0.395 | 0.4 |  | 0.405 | 0.36826 |  | 0.45 |  | 0.31 | 0.425 | 0.4 | 0.42 | 0.415 | 0.4
| Std Dev. |  | 0.017607 | 0.01472 | 0.01681 | 0.06957 |  | 0.018708 | 0.016432 | 0.01472 | 0.021602 | 0.01472 | 0.0083666 |  | 0.008165 | 0.015822 |  | 0.020412 |  | 0.0083666 | 0.020976 | 0.018348 | 0.013292 | 0.01472 | 0.016432
| Rel.Std.Dev. |  | 4.457% | 3.891% | 3.881% | 16.97% |  | 4.736% | 4.057% | 3.605% | 5.356% | 3.758% | 2.066% |  | 2.024% | 4.281% |  | 4.519% |  | 2.743% | 4.878% | 4.457% | 3.177% | 3.576% | 4.16%
| PDM3 |  | -1.486% | -5.643% | 8.036% | 2.255% |  | -1.486% | 1.008% | 1.839% | 0.5922% | -2.317% | 1.008% |  | 0.5922% | -7.816% |  | 12.65% |  | -23.93% | 7.243% | 2.671% | 4.333% | 2.671% | -1.486%
| Z-Score (Absolute) |  | 0.51 | 1.35 | 1.43 | 0.25 | 4.8 | 0.51 | 0 | 0.17 | 0.08 | 0.67 | 0 | 7.84 | 0.08 | 1.79 | 4.8 | 2.36 | 4.8 | 5.06 | 1.26 | 0.34 | 0.67 | 0.34 | 0.51
| Analytical results for Ce in OREAS 906b (Certified Value 77 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 0
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [6 empty] | 1
| 1 | 1 | 76.511 | 80.3 | 73.4173 | 70.1 | 72.7 | 73.7 | 81.7 | 74.8 | 71.5 | 64.46 | 77.51 | 69.4384 | 73.8 | 84.3864 | 74 | 78.9 | 91.79 | 73 | 81.5 | 80.8 | [6 empty] | 1
| 1 | 2 | 76.96 | 80.9 | 74.0676 | 69.1 | 74.4 | 74.1 | 80.8 | 77 | 72.2 | 66.6 | 75.82 | 68.7121 | 71.4 | 84.199 | 74 | 79.8 | 83.27 | 77 | 83.7 | 80.5 | [6 empty] | 39
| 1 | 3 | 75.209 | 80.7 | 78.6238 | 72.2 | 76.9 | 75.1 | 80.9 | 75.7 | 72.2 | 66.23 | 78.46 | 69.6094 | 71.3 | 84.112 | 72 | 81.3 | 87.52 | 77 | 82.2 | 81.6 | [6 empty] | 16
| 1 | 4 | 78.024 | 81 | 75.9773 | 69.8 | 73.5 | 79.4 | 80.5 | 74.6 | 70.3 | 67.1 | 74.9 | 69.212 | 73.3 | 84.084 | 74 | 76.8 | 91.58 | 75 | 82.8 | 79.6 | [6 empty] | 76.5835
| 1 | 5 | 76.285 | 80.5 | 80.4119 | 70.7 | 75.7 | 75.6 | 85.4 | 74.1 | 73.8 | 67.08 | 74.66 | 69.8641 | 72.9 | 84.052 | 73 | 82.8 | 85.37 | 79 | 83 | 78.6 | [6 empty] | 134
| 1 | 6 | 75.899 | 80.6 | 77.0306 | 71.1 | 73.2 | 73.4 | 83.1 | 76.9 | 71.1 | 67.91 | 76.32 | 68.7772 | 74 | 84.1744 | 72 | 83 | 88.52 | 77 | 81.3 | 81.7
| Mean |  | 76.4813 | 80.6667 | 76.5881 | 70.5 | 74.4 | 75.2167 | 82.0667 | 75.5167 | 71.85 | 66.5633 | 76.2783 | 69.2689 | 72.7833 | 84.168 | 73.1667 | 80.4333 | 88.0083 | 76.3333 | 82.4167 | 80.4667
| Median |  | 76.398 | 80.65 | 76.5039 | 70.4 | 73.95 | 74.6 | 81.3 | 75.25 | 71.85 | 66.84 | 76.07 | 69.3252 | 73.1 | 84.1432 | 73.5 | 80.55 | 88.02 | 77 | 82.5 | 80.65
| Std Dev. |  | 0.95982 | 0.2582 | 2.6736 | 1.0863 | 1.6174 | 2.214 | 1.8833 | 1.2254 | 1.1946 | 1.1752 | 1.4837 | 0.45915 | 1.1754 | 0.12029 | 0.98319 | 2.4039 | 3.3755 | 2.0656 | 0.92394 | 1.1961
| Rel.Std.Dev. |  | 1.255% | 0.3201% | 3.491% | 1.541% | 2.174% | 2.943% | 2.295% | 1.623% | 1.663% | 1.765% | 1.945% | 0.6628% | 1.615% | 0.1429% | 1.344% | 2.989% | 3.835% | 2.706% | 1.121% | 1.486%
| PDM3 |  | -0.1334% | 5.332% | 0.006002% | -7.944% | -2.851% | -1.785% | 7.16% | -1.393% | -6.181% | -13.08% | -0.3984% | -9.551% | -4.962% | 9.904% | -4.462% | 5.027% | 14.92% | -0.3266% | 7.617% | 5.071%
| Z-Score (Absolute) |  | 0.03 | 0.71 | 0.05 | 0.94 | 0.31 | 0.18 | 0.94 | 0.13 | 0.73 | 1.59 | 0 | 1.15 | 0.57 | 1.28 | 0.51 | 0.67 | 1.9 | 0 | 0.99 | 0.68
| Analytical results for Co in OREAS 906b (Certified Value 15.8 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 2
| 1 | 1 | 16.2 | 16.4 | 15.1583 | 14 | 13.7 | 14.6 | 16 | 15.2 | 14.5 | 15.6 | 15.1 | 15.1 | 15.5 | 15.8049 | 16 | 15.44 | 16.023 | 16 | 16.3 | 16.1 | 22.61 | 16 | 16.7 | 16.6 |  |  | 1
| 1 | 2 | 16.2 | 16.3 | 15.5719 | 16 | 13 | 16.2 | 16.3 | 15.4 | 14.8 | 15.8 | 15.4 | 15 | 15.9 | 15.9293 | 16 | 15.48 | 16.3603 | 17 | 15.7 | 15.8 | 23.495 | 16.2 | 16.8 | 16.3 |  |  | 27
| 1 | 3 | 16.2 | 16.4 | 16.3979 | 15.4 | 13.6 | 17 | 16.3 | 15.2 | 14.8 | 16 | 15.2 | 15 | 15.5 | 16.0125 | 16 | 14.89 | 16.1008 | 16 | 17.1 | 16.3 | 22.94 | 15.7 | 16.3 | 16.4 |  |  | 16
| 1 | 4 | 15.9 | 16.2 | 15.213 | 13.7 | 13.3 | 19.4 | 16.4 | 15 | 14.5 | 15.6 | 15.8 | 15.5 | 15.5 | 16.0461 | 17 | 15.36 | 16.235 | 16 | 15.6 | 16.4 | 22.015 | 16 | 16.1 | 16.2 |  |  | 15.8359
| 1 | 5 | 15.9 | 16.4 | 16.7124 | 15 | 13.8 | 15.7 | 16.2 | 15.7 | 14.6 | 16.2 | 14.7 | 16 | 15.9 | 16.4651 | 16 | 15.21 | 16.1558 | 17 | 16.8 | 16.6 | 23.235 | 16.3 | 16.4 | 16.4 |  |  | 135
| 1 | 6 | 16.3 | 16.3 | 16.0216 | 14.6 | 13.6 | 15.6 | 16.4 | 15.8 | 14.8 | 15.3 | 15.5 | 16 | 15.6 | 15.6801 | 16 | 15.43 | 16.2187 | 17 | 16.3 | 17.1 | 22.57 | 16.1 | 16.7 | 15.6
| Mean |  | 16.1167 | 16.3333 | 15.8458 | 14.7833 | 13.5 | 16.4167 | 16.2667 | 15.3833 | 14.6667 | 15.75 | 15.2833 | 15.4333 | 15.65 | 15.9897 | 16.1667 | 15.3017 | 16.1823 | 16.5 | 16.3 | 16.3833 | 22.8108 | 16.05 | 16.5 | 16.25
| Median |  | 16.2 | 16.35 | 15.7967 | 14.8 | 13.6 | 15.95 | 16.3 | 15.3 | 14.7 | 15.7 | 15.3 | 15.3 | 15.55 | 15.9709 | 16 | 15.395 | 16.1873 | 16.5 | 16.3 | 16.35 | 22.775 | 16.05 | 16.55 | 16.35
| Std Dev. |  | 0.17224 | 0.08165 | 0.63793 | 0.86352 | 0.29665 | 1.6594 | 0.15055 | 0.31252 | 0.15055 | 0.32094 | 0.37639 | 0.4761 | 0.19748 | 0.26959 | 0.40825 | 0.22302 | 0.11705 | 0.54772 | 0.58992 | 0.4446 | 0.52846 | 0.20736 | 0.27568 | 0.34496
| Rel.Std.Dev. |  | 1.069% | 0.4999% | 4.026% | 5.841% | 2.197% | 10.11% | 0.9255% | 2.032% | 1.027% | 2.038% | 2.463% | 3.085% | 1.262% | 1.686% | 2.525% | 1.457% | 0.7233% | 3.32% | 3.619% | 2.714% | 2.317% | 1.292% | 1.671% | 2.123%
| PDM3 |  | 1.773% | 3.141% | 0.06261% | -6.647% | -14.75% | 3.667% | 2.72% | -2.858% | -7.384% | -0.5426% | -3.489% | -2.542% | -1.174% | 0.9709% | 2.089% | -3.374% | 2.187% | 4.193% | 2.931% | 3.457% | 44.04% | 1.352% | 4.193% | 2.615%
| Z-Score (Absolute) |  | 0.18 | 0.58 | 0.32 | 2.27 | 4.63 | 0.73 | 0.45 | 1.17 | 2.49 | 0.5 | 1.35 | 1.08 | 0.68 | 0.06 | NA | 1.32 | 0.3 | NA | 0.52 | 0.67 | 12.49 | 0.06 | 0.88 | 0.42
Results from laboratories 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 906b (Certified Value 10 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  |  | 2
| 1 | 1 | 10 | 10 | 9.3957 | 9 | 10 | 10 | 10 | 9 | 10 | 10 | 11.7 | 10 | 10.2831 | 9 | 11 | 10 | 10 | 9 | 8 | 10 | 12 | 10.9 |  |  |  |  | 1
| 1 | 2 | 11 | 10 | 10.3727 | 10.1 | 11 | 10 | 10 | 10 | 10 | 11 | 12 | 9 | 10.887 | 9 | 11.3 | 11 | 9 | 9 | 8 | 10 | 12 | 10.5 |  |  |  |  | 13
| 1 | 3 | 10 | 9 | 9.4529 | 10.5 | 10 | 9 | 10 | 9 | 10 | 11 | 13 | 10 | 10.1694 | 9 | 10.7 | 10 | 8 | 9 | 8 | 11 | 9 | 10.5 |  |  |  |  | 16
| 1 | 4 | 10 | 10 | 10.0457 | 9.8 | 11 | 10 | 11 | 9 | 9 | 10 | 12.7 | 10 | 11.0901 | 9 | 10.3 | 10 | 7 | 10 | 8 | 11 | 10 | 10.5 |  |  |  |  | 10.029
| 1 | 5 | 10 | 10 | 9.8875 | 9.7 | 10 | 11 | 11 | 9 | 10 | 9 | 12 | 10 | 10.47 | 10 | 10.7 | 10 | 11 | 10 | 8 | 11 | 10 | 10.7 |  |  |  |  | 136
| 1 | 6 | 10 | 10 | 10.7088 | 9.8 | 10 | 10 | 11 | 10 | 10 | 10 | 12 | 9 | 10.1195 | 8 | 10.7 | 11 | 8 | 9 | 8 | 11 | 10 | 10.7
| Mean |  | 10.1667 | 9.8333 | 9.9772 | 9.8167 | 10.3333 | 10 | 10.5 | 9.3333 | 9.8333 | 10.1667 | 12.2333 | 9.6667 | 10.5032 | 9 | 10.7833 | 10.3333 | 8.8333 | 9.3333 | 8 | 10.6667 | 10.5 | 10.6333
| Median |  | 10 | 10 | 9.9666 | 9.8 | 10 | 10 | 10.5 | 9 | 10 | 10 | 12 | 10 | 10.3766 | 9 | 10.7 | 10 | 8.5 | 9 | 8 | 11 | 10 | 10.6
| Std Dev. |  | 0.40825 | 0.40825 | 0.51331 | 0.49565 | 0.5164 | 0.63246 | 0.54772 | 0.5164 | 0.40825 | 0.75277 | 0.50067 | 0.5164 | 0.40001 | 0.63246 | 0.33714 | 0.5164 | 1.472 | 0.5164 | 0 | 0.5164 | 1.2247 | 0.1633
| Rel.Std.Dev. |  | 4.016% | 4.152% | 5.145% | 5.049% | 4.997% | 6.325% | 5.216% | 5.533% | 4.152% | 7.404% | 4.093% | 5.342% | 3.808% | 7.027% | 3.127% | 4.997% | 16.66% | 5.533% | 0% | 4.841% | 11.66% | 1.536%
| PDM3 |  | 1.372% | -1.951% | -0.5166% | -2.117% | 3.034% | -0.2894% | 4.696% | -6.937% | -1.951% | 1.372% | 21.98% | -3.613% | 4.728% | -10.26% | 7.521% | 3.034% | -11.92% | -6.937% | -20.23% | 6.358% | 4.696% | 6.026%
| Z-Score (Absolute) |  | 0.13 | 0.4 | 0.17 | 0.43 | 0.4 | 0.13 | 0.67 | 1.21 | 0.4 | 0.13 | 3.48 | 0.67 | 0.68 | 1.75 | 1.13 | 0.4 | 2.02 | 1.21 | 3.37 | 0.94 | 0.67 | 0.89
| Analytical results for Cs in OREAS 906b (Certified Value 0.98 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.5g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [9 empty] | 2
| 1 | 1 | 1 | 1.24 | 1.008 | 0.81 | 0.81 | 1.05 | 0.76 | 0.78 | 0.87 | 1.0311 | 0.92 | 1.0183 | 2.1 | 1.29 | 0.89 | 0.93 | 1.08 | [9 empty] | 1
| 1 | 2 | 0.97 | 1.26 | 1.0088 | 0.83 | 0.83 | 1 | 0.83 | 0.79 | 0.87 | 0.99794 | 0.93 | 1.144 | 2 | 1.2 | 0.97 | 0.94 | 1.14 | [9 empty] | 42
| 1 | 3 | 0.97 | 1.28 | 1.0475 | 0.84 | 0.82 | 1.04 | 0.78 | 0.79 | 0.86 | 1.0532 | 0.92 | 1.1571 | 2 | < 0.05 | 0.92 | 0.93 | 1.14 | [9 empty] | 16
| 1 | 4 | 1.01 | 1.29 | 1.0082 | 0.82 | 0.84 | 1.01 | 0.77 | 0.78 | 0.85 | 1.0358 | 0.91 | 1.0687 | 2 | 1.24 | 0.94 | 0.93 | 1.07 | [9 empty] | 0.97597
| 1 | 5 | 0.94 | 1.25 | 1.0499 | 0.85 | 0.82 | 1.06 | 0.78 | 0.79 | 0.87 | 1.0327 | 0.95 | 1.0402 | 1.9 | 1.24 | 0.94 | 0.95 | 1.06 | [9 empty] | 137
| 1 | 6 | 0.98 | 1.28 | 1.0045 | 0.78 | 0.79 | 1.04 | 0.77 | 0.79 | 0.84 | 0.99354 | 0.91 | 1.0798 | 2.1 | 1.29 | 0.93 | 0.94 | 1.14
| Mean |  | 0.97833 | 1.2667 | 1.0211 | 0.82167 | 0.81833 | 1.0333 | 0.78167 | 0.78667 | 0.86 | 1.024 | 0.92333 | 1.0847 | 2.0167 | 1.252 | 0.93167 | 0.93667 | 1.105
| Median |  | 0.975 | 1.27 | 1.0085 | 0.825 | 0.82 | 1.04 | 0.775 | 0.79 | 0.865 | 1.0319 | 0.92 | 1.0743 | 2 | 1.24 | 0.935 | 0.935 | 1.11
| Std Dev. |  | 0.024833 | 0.019664 | 0.021403 | 0.024833 | 0.017224 | 0.023381 | 0.024833 | 0.005164 | 0.012649 | 0.023342 | 0.015055 | 0.055534 | 0.075277 | 0.038341 | 0.026394 | 0.008165 | 0.038859
| Rel.Std.Dev. |  | 2.538% | 1.552% | 2.096% | 3.022% | 2.105% | 2.263% | 3.177% | 0.6564% | 1.471% | 2.279% | 1.631% | 5.12% | 3.733% | 3.062% | 2.833% | 0.8717% | 3.517%
| PDM3 |  | 0.2421% | 29.79% | 4.628% | -15.81% | -16.15% | 5.877% | -19.91% | -19.4% | -11.88% | 4.926% | -5.393% | 11.14% | 106.6% | 28.28% | -4.539% | -4.027% | 13.22%
| Z-Score (Absolute) |  | 0.15 | 2.22 | 0.46 | 0.98 | 1 | 0.55 | 1.26 | 1.23 | 0.7 | 0.48 | 0.25 | 0.91 | NA | 0.65 | 0.19 | 0.15 | 1.06
Results from laboratory 2.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 906b (Certified Value 0.316 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 3
| 1 | 1 | 0.32627 | 0.298 | 0.30856 | 0.29374 | 0.3148 | 0.315 | 0.325 | 0.334 | 0.313 | 0.306 | 0.3054 | 0.326 | 0.31255 | 0.32336 | 0.3034 | 0.322 | 0.29527 | 0.3118 | 0.331 | 0.305 | 0.3234 | 0.31673 | 0.3143 | 0.315 | 0.3567 | 0.32133 | 1
| 1 | 2 | 0.32704 | 0.3 | 0.30816 | 0.30861 | 0.3198 | 0.323 | 0.322 | 0.329 | 0.322 | 0.304 | 0.3119 | 0.333 | 0.31744 | 0.32169 | 0.3089 | 0.3168 | 0.29142 | 0.3123 | 0.338 | 0.317 | 0.3241 | 0.31605 | 0.316 | 0.314 | 0.3556 | 0.31838 | 29
| 1 | 3 | 0.3222 | 0.298 | 0.30888 | 0.31205 | 0.3165 | 0.33 | 0.322 | 0.322 | 0.319 | 0.305 | 0.3079 | 0.329 | 0.31637 | 0.32276 | 0.306 | 0.3156 | 0.29086 | 0.313 | 0.35 | 0.318 | 0.3229 | 0.31793 | 0.3152 | 0.315 | 0.346 | 0.3213 | 16
| 1 | 4 | 0.33301 | 0.303 | 0.30973 | 0.30358 | 0.3214 | 0.331 | 0.307 | 0.328 | 0.322 | 0.302 | 0.3128 | 0.338 | 0.31629 | 0.32451 | 0.3105 | 0.3222 | 0.29473 | 0.3112 | 0.325 | 0.3084 | 0.3245 | 0.31594 | 0.3152 | 0.309 | 0.3514 | 0.32244 | 0.31608
| 1 | 5 | 0.32167 | 0.3 | 0.30974 | 0.30262 | 0.3224 | 0.328 | 0.326 | 0.332 | 0.315 | 0.299 | 0.3006 | 0.333 | 0.31219 | 0.32611 | 0.3055 | 0.3184 | 0.29505 | 0.3135 | 0.344 | 0.3236 | 0.3241 | 0.31541 | 0.3152 | 0.315 | 0.3539 | 0.31847 | 138
| 1 | 6 | 0.32522 | 0.298 | 0.31011 | 0.30389 | 0.3129 | 0.318 | 0.315 | 0.331 | 0.323 | 0.305 | 0.3061 | 0.339 | 0.31031 | 0.32348 | 0.3107 | 0.3219 | 0.2906 | 0.3154 | 0.338 | 0.3169 | 0.3223 | 0.31604 | 0.3101 | 0.316 | 0.3513 | 0.32164
| Mean |  | 0.3259 | 0.2995 | 0.30919 | 0.30408 | 0.31797 | 0.32417 | 0.3195 | 0.32933 | 0.319 | 0.3035 | 0.30745 | 0.333 | 0.31419 | 0.32365 | 0.3075 | 0.31948 | 0.29299 | 0.31287 | 0.33767 | 0.31482 | 0.32355 | 0.31635 | 0.31433 | 0.314 | 0.35248 | 0.32059
| Median |  | 0.32574 | 0.299 | 0.3093 | 0.30374 | 0.31815 | 0.3255 | 0.322 | 0.33 | 0.3205 | 0.3045 | 0.307 | 0.333 | 0.31442 | 0.32342 | 0.30745 | 0.32015 | 0.29308 | 0.31265 | 0.338 | 0.31695 | 0.32375 | 0.31605 | 0.3152 | 0.315 | 0.35265 | 0.32132
| Std Dev. |  | 0.0040996 | 0.0019748 | 0.00077444 | 0.0062184 | 0.0038119 | 0.0066156 | 0.0072319 | 0.0041793 | 0.0041473 | 0.0025884 | 0.0045072 | 0.00502 | 0.0028799 | 0.0015171 | 0.0029752 | 0.0029329 | 0.0022426 | 0.0014882 | 0.0089144 | 0.00684 | 0.00083845 | 0.00088075 | 0.0021426 | 0.0025298 | 0.0038499 | 0.0017309
| Rel.Std.Dev. |  | 1.258% | 0.6594% | 0.2505% | 2.045% | 1.199% | 2.041% | 2.263% | 1.269% | 1.3% | 0.8529% | 1.466% | 1.507% | 0.9166% | 0.4688% | 0.9676% | 0.918% | 0.7654% | 0.4757% | 2.64% | 2.173% | 0.2591% | 0.2784% | 0.6816% | 0.8057% | 1.092% | 0.5399%
| PDM3 |  | 3.106% | -5.247% | -2.18% | -3.797% | 0.5954% | 2.557% | 1.08% | 4.191% | 0.9223% | -3.981% | -2.732% | 5.351% | -0.5989% | 2.393% | -2.716% | 1.075% | -7.307% | -1.018% | 6.828% | -0.4012% | 2.362% | 0.0839% | -0.5541% | -0.6596% | 11.52% | 1.427%
| Z-Score (Absolute) |  | 0.87 | 1.76 | 0.8 | 1.31 | 0.08 | 0.7 | 0.23 | 1.22 | 0.18 | 1.36 | 0.97 | 1.58 | 0.3 | 0.65 | 0.96 | 0.23 | 2.41 | 0.43 | 2.05 | 0.23 | 0.64 | 0.08 | 0.28 | 0.32 | 3.53 | 0.34
| Analytical results for Dy in OREAS 906b (Indicative Value 1.8 ppm) | [27 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 1.671 | 1.5955 | 2.0563 | 1.9 | [22 empty] | 1
| 1 | 2 | 1.689 | 1.6535 | 2.0247 | 1.8 | [22 empty] | 16
| 1 | 3 | 1.675 | 1.6781 | 2.0321 | 1.9 | [22 empty] | 16
| 1 | 4 | 1.735 | 1.6089 | 1.9801 | 1.8 | [22 empty] | 1.7998
| 1 | 5 | 1.676 | 1.6643 | 1.9652 | 1.9 | [22 empty] | 22
| 1 | 6 | 1.678 | 1.597 | 1.9165 | 2
| Mean |  | 1.6873 | 1.6329 | 1.9958 | 1.8833
| Median |  | 1.677 | 1.6312 | 2.0024 | 1.9
| Std Dev. |  | 0.024122 | 0.036644 | 0.051517 | 0.075277
| Rel.Std.Dev. |  | 1.43% | 2.244% | 2.581% | 3.997%
| PDM3 |  | -6.251% | -9.276% | 10.89% | 4.639%
| Z-Score (Absolute) |  | 0.53 | 0.82 | 1.13 | 0.53
| Analytical results for Er in OREAS 906b (Indicative Value 0.47 ppm) | [27 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 0.419 | 0.39036 | 0.60215 | 0.5 | [22 empty] | 1
| 1 | 2 | 0.428 | 0.4064 | 0.58477 | 0.5 | [22 empty] | 17
| 1 | 3 | 0.423 | 0.42496 | 0.57095 | 0.5 | [22 empty] | 16
| 1 | 4 | 0.432 | 0.38431 | 0.60623 | 0.5 | [22 empty] | 0.47267
| 1 | 5 | 0.411 | 0.40387 | 0.55301 | 0.4 | [22 empty] | 23
| 1 | 6 | 0.431 | 0.42107 | 0.55192 | 0.5
| Mean |  | 0.424 | 0.40516 | 0.57817 | 0.48333
| Median |  | 0.4255 | 0.40514 | 0.57786 | 0.5
| Std Dev. |  | 0.0080498 | 0.016137 | 0.023579 | 0.040825
| Rel.Std.Dev. |  | 1.899% | 3.983% | 4.078% | 8.447%
| PDM3 |  | -10.3% | -14.28% | 22.32% | 2.257%
| Z-Score (Absolute) |  | 0.51 | 0.84 | 2.15 | 0.51
| Analytical results for Eu in OREAS 906b (Indicative Value 0.99 ppm) | [27 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 0.956 | 0.85167 | 1.1269 | 1 | [22 empty] | 1
| 1 | 2 | 0.937 | 0.84614 | 1.1632 | 1 | [22 empty] | 18
| 1 | 3 | 0.935 | 0.86113 | 1.1708 | 1 | [22 empty] | 16
| 1 | 4 | 0.952 | 0.86533 | 1.1432 | 1 | [22 empty] | 0.98551
| 1 | 5 | 0.949 | 0.88233 | 1.1235 | 1 | [22 empty] | 24
| 1 | 6 | 0.949 | 0.85746 | 1.0825 | 1
| Mean |  | 0.94633 | 0.86068 | 1.135 | 1
| Median |  | 0.949 | 0.8593 | 1.1351 | 1
| Std Dev. |  | 0.0084301 | 0.012599 | 0.031921 | 0
| Rel.Std.Dev. |  | 0.8908% | 1.464% | 2.812% | 0%
| PDM3 |  | -3.975% | -12.67% | 15.17% | 1.471%
| Z-Score (Absolute) |  | 0.26 | 1.09 | 1.57 | 0.26
| Analytical results for Fe in OREAS 906b (Certified Value 4.22 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 4.17 | 4.22 | 4.2544 | 3.82 | 4.34 | 4.19 | 4.25 | 4.41 | 4.12 | 4.1 | 4.02 | 4.47 | 4.21 | 4.2052 | 4.27 | 4.18 | 3.8 | 4.3 | 3.91 | 4.17 | 4.9062 | 4.35 | 4.43 | 3.9145 | 4.16 |  | 1
| 1 | 2 | 4.31 | 4.18 | 4.4334 | 3.95 | 4.31 | 4.24 | 4.24 | 4.32 | 4.21 | 4.07 | 4.16 | 4.36 | 4.2 | 4.1522 | 4.26 | 4.12 | 3.74 | 3.94 | 4.07 | 4.15 | 4.8602 | 4.33 | 4.37 | 3.8254 | 4.14 |  | #N/A
| 1 | 3 | 4.3 | 4.23 | 4.4171 | 4.05 | 4.33 | 4.31 | 4.21 | 4.27 | 4.18 | 4.06 | 4.28 | 4.36 | 4.25 | 4.1728 | 4.32 | 4.11 | 3.6 | 4.1 | 4.14 | 4.15 | 4.7925 | 4.37 | 4.38 | 3.7266 | 4.12 |  | 16
| 1 | 4 | 4.31 | 4.28 | 4.2749 | 3.98 | 4.3 | 4.3 | 4.19 | 4.33 | 4.25 | 4.04 | 4.14 | 4.36 | 4.19 | 4.1918 | 4.34 | 4.17 | 3.69 | 3.9 | 3.99 | 4.28 | 4.7794 | 4.3 | 4.31 | 3.836 | 4.14 |  | 4.2176
| 1 | 5 | 4.18 | 4.24 | 4.386 | 3.93 | 4.32 | 4.32 | 4.24 | 4.39 | 4.16 | 4.06 | 3.99 | 4.44 | 4.15 | 4.1702 | 4.29 | 4.1 | 3.72 | 4.77 | 4.15 | 4.26 | 4.8427 | 4.37 | 4.36 | 3.8666 | 4.11 |  | 139
| 1 | 6 | 4.24 | 4.24 | 4.2985 | 3.92 | 4.3 | 4.22 | 4.19 | 4.37 | 4.24 | 4.08 | 4.12 | 4.61 | 4.28 | 4.1715 | 4.3 | 4.12 | 3.52 | 4.06 | 4.11 | 4.46 | 4.8138 | 4.39 | 4.47 | 3.786 | 4.11
| Mean |  | 4.2517 | 4.2317 | 4.344 | 3.9417 | 4.3167 | 4.2633 | 4.22 | 4.3483 | 4.1933 | 4.0683 | 4.1183 | 4.4333 | 4.2133 | 4.1773 | 4.2967 | 4.1333 | 3.6783 | 4.1783 | 4.0617 | 4.245 | 4.8325 | 4.3517 | 4.3867 | 3.8259 | 4.13
| Median |  | 4.27 | 4.235 | 4.3422 | 3.94 | 4.315 | 4.27 | 4.225 | 4.35 | 4.195 | 4.065 | 4.13 | 4.4 | 4.205 | 4.1722 | 4.295 | 4.12 | 3.705 | 4.08 | 4.09 | 4.215 | 4.8282 | 4.36 | 4.375 | 3.8307 | 4.13
| Std Dev. |  | 0.064936 | 0.032506 | 0.077434 | 0.075741 | 0.01633 | 0.053914 | 0.026833 | 0.051543 | 0.049666 | 0.020412 | 0.10439 | 0.098725 | 0.045898 | 0.018547 | 0.030111 | 0.033267 | 0.10167 | 0.32227 | 0.094322 | 0.11946 | 0.047049 | 0.032506 | 0.056095 | 0.064902 | 0.02
| Rel.Std.Dev. |  | 1.527% | 0.7682% | 1.783% | 1.922% | 0.3783% | 1.265% | 0.6358% | 1.185% | 1.184% | 0.5017% | 2.535% | 2.227% | 1.089% | 0.444% | 0.7008% | 0.8048% | 2.764% | 7.713% | 2.322% | 2.814% | 0.9736% | 0.747% | 1.279% | 1.696% | 0.4843%
| PDM3 |  | 0.8087% | 0.3345% | 2.999% | -6.542% | 2.35% | 1.085% | 0.05783% | 3.101% | -0.5744% | -3.538% | -2.353% | 5.116% | -0.1002% | -0.9548% | 1.876% | -1.997% | -12.79% | -0.9301% | -3.696% | 0.6506% | 14.58% | 3.18% | 4.01% | -9.288% | -2.076%
| Z-Score (Absolute) |  | 0.22 | 0.08 | 0.87 | 1.94 | 0.67 | 0.3 | 0 | 0.9 | 0.19 | 1.06 | 0.71 | 1.49 | 0.05 | 0.3 | 0.53 | 0.6 | 3.78 | 0.29 | 1.1 | 0.17 | 4.27 | 0.92 | 1.16 | 2.75 | 0.63
| Analytical results for Ga in OREAS 906b (Certified Value 7 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 6.86 | 7.4 | 6.6885 | 6.6 | 12 | 6.52 | 6.86 | 6.9 | 6.46 | 6.94 | 6.9 | < 10 | 6.6 | 7.9354 | 7.39 | 7 | 10.1 | 8.8 | 6.94 | 9 | 7 | 7.28 | 6.7 |  |  |  | 1
| 1 | 2 | 6.81 | 7.4 | 6.7947 | 7.1 | 11.7 | 6.75 | 7 | 6.87 | 6.73 | 7.05 | 7 | < 10 | 7 | 7.9281 | 6.89 | 7 | 10.1 | 8.1 | 7.11 | 9 | 7.1 | 7.57 | 6.5 |  |  |  | 47
| 1 | 3 | 7.01 | 7.4 | 7.0202 | 6.7 | 12.1 | 7.08 | 6.87 | 6.79 | 6.56 | 7.01 | 7.1 | < 10 | 6.6 | 7.9813 | 6.91 | 6 | 10.7 | 8.8 | 6.85 | 9 | 7.1 | 7.47 | 6.6 |  |  |  | 16
| 1 | 4 | 7.12 | 7.6 | 7.1187 | 6.4 | 11.6 | 6.77 | 6.47 | 6.69 | 6.48 | 6.86 | 7.3 | < 10 | 6.5 | 8.0067 | 7.26 | 7 | 9.73 | 8.9 | 7 | 9 | 6.9 | 7.13 | 6.6 |  |  |  | 6.9987
| 1 | 5 | 6.52 | 7.5 | 7.1435 | 6.8 | 11.4 | 6.7 | 6.98 | 7.08 | 6.44 | 6.94 | 7 | < 10 | 6.8 | 8.5561 | 7.22 | 7 | 9.8 | 8.6 | 7.23 | 9 | 7.1 | 7.25 | 6.6 |  |  |  | 140
| 1 | 6 | 7.06 | 7.5 | 7.269 | 6.8 | 12.1 | 6.33 | 6.79 | 6.97 | 6.66 | 6.79 | 7.2 | < 10 | 6.7 | 7.8033 | 7.4 | 6 | 10.9 | 8.5 | 6.98 | 9 | 7.1 | 7.42 | 6.6
| Mean |  | 6.8967 | 7.4667 | 7.0058 | 6.7333 | 11.8167 | 6.6917 | 6.8283 | 6.8833 | 6.555 | 6.9317 | 7.0833 |  | 6.7 | 8.0352 | 7.1783 | 6.6667 | 10.2217 | 8.6167 | 7.0183 | 9 | 7.05 | 7.3533 | 6.6
| Median |  | 6.935 | 7.45 | 7.0695 | 6.75 | 11.85 | 6.725 | 6.865 | 6.885 | 6.52 | 6.94 | 7.05 |  | 6.65 | 7.9583 | 7.24 | 7 | 10.1 | 8.7 | 6.99 | 9 | 7.1 | 7.35 | 6.6
| Std Dev. |  | 0.21915 | 0.08165 | 0.22201 | 0.23381 | 0.29269 | 0.25325 | 0.1924 | 0.13619 | 0.11794 | 0.095376 | 0.1472 |  | 0.17889 | 0.26467 | 0.22693 | 0.5164 | 0.47709 | 0.29269 | 0.13378 | 0 | 0.083666 | 0.16182 | 0.063246
| Rel.Std.Dev. |  | 3.178% | 1.094% | 3.169% | 3.472% | 2.477% | 3.785% | 2.818% | 1.978% | 1.799% | 1.376% | 2.078% |  | 2.67% | 3.294% | 3.161% | 7.746% | 4.667% | 3.397% | 1.906% | 0% | 1.187% | 2.201% | 0.9583%
| PDM3 |  | -1.458% | 6.686% | 0.1006% | -3.792% | 68.84% | -4.387% | -2.435% | -1.649% | -6.34% | -0.9582% | 1.209% |  | -4.268% | 14.81% | 2.566% | -4.745% | 46.05% | 23.12% | 0.2801% | 28.59% | 0.7326% | 5.067% | -5.697%
| Z-Score (Absolute) |  | 0.24 | 1.02 | 0 | 0.6 | 10.61 | 0.69 | 0.39 | 0.27 | 0.99 | 0.16 | 0.17 | 4.42 | 0.67 | 2.27 | 0.38 | NA | 7.09 | 3.55 | 0.03 | NA | 0.1 | 0.77 | 0.89
Results from laboratory 2.19 and 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 906b (Indicative Value 4.07 ppm) | [27 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 3.554 | 3.5944 | 5.2874 | 4.1 | [22 empty] | 1
| 1 | 2 | 3.557 | 3.5257 | 5.0543 | 3.9 | [22 empty] | 19
| 1 | 3 | 3.5 | 3.6193 | 5.2349 | 4 | [22 empty] | 16
| 1 | 4 | 3.657 | 3.5773 | 5.1236 | 4 | [22 empty] | 4.0711
| 1 | 5 | 3.572 | 3.5712 | 5.0308 | 4 | [22 empty] | 25
| 1 | 6 | 3.557 | 3.5329 | 4.9577 | 4.2
| Mean |  | 3.5662 | 3.5701 | 5.1148 | 4.0333
| Median |  | 3.557 | 3.5742 | 5.0889 | 4
| Std Dev. |  | 0.050949 | 0.035832 | 0.12625 | 0.10328
| Rel.Std.Dev. |  | 1.429% | 1.004% | 2.468% | 2.561%
| PDM3 |  | -12.4% | -12.31% | 25.64% | -0.9276%
| Z-Score (Absolute) |  | 0.68 | 0.67 | 3.79 | 0.67
| Analytical results for Ge in OREAS 906b (Certified Value 0.11 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.20 | 2.22 | 2.24 | 2.25 | [14 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [14 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.1g | 0.5g | 0.25g | 0.5g | [14 empty] | 2
| 1 | 1 | < 5 | 0.08 | 0.11 | 0.12 | 0.08 | 0.1 | 0.1 | 0.26848 | 0.15 | 0.35 | < 0.1 | < 0.1 | [14 empty] | 1
| 1 | 2 | < 5 | 0.07 | 0.11 | 0.12 | 0.09 | 0.12 | 0.1 | 0.26686 | 0.15 | 0.33 | < 0.1 | < 0.1 | [14 empty] | 49
| 1 | 3 | < 5 | 0.09 | 0.12 | 0.11 | 0.08 | 0.1 | 0.1 | 0.28366 | 0.16 | 0.34 | < 0.1 | < 0.1 | [14 empty] | 16
| 1 | 4 | < 5 | 0.08 | 0.11 | 0.11 | 0.08 | 0.11 | 0.1 | 0.27373 | 0.15 | 0.35 | < 0.1 | < 0.1 | [14 empty] | 0.10833
| 1 | 5 | < 5 | 0.07 | 0.11 | 0.11 | 0.08 | 0.12 | 0.1 | 0.28519 | 0.16 | 0.32 | < 0.1 | < 0.1 | [14 empty] | 141
| 1 | 6 | < 5 | 0.07 | 0.11 | 0.12 | 0.1 | 0.09 | 0.1 | 0.27453 | 0.16 | 0.34 | < 0.1 | < 0.1
| Mean |  |  | 0.076667 | 0.11167 | 0.115 | 0.085 | 0.10667 | 0.1 | 0.27541 | 0.155 | 0.33833
| Median |  |  | 0.075 | 0.11 | 0.115 | 0.08 | 0.105 | 0.1 | 0.27413 | 0.155 | 0.34
| Std Dev. |  |  | 0.008165 | 0.0040825 | 0.0054772 | 0.0083666 | 0.012111 | 1.5202e-17 | 0.0075949 | 0.0054772 | 0.01169
| Rel.Std.Dev. |  |  | 10.65% | 3.656% | 4.763% | 9.843% | 11.35% | 1.52e-14% | 2.758% | 3.534% | 3.455%
| PDM3 |  |  | -29.23% | 3.077% | 6.154% | -21.54% | -1.538% | -7.692% | 154.2% | 43.08% | 212.3%
| Z-Score (Absolute) |  | 37.16 | 0.54 | 0 | 0.05 | 0.41 | 0.08 | NA | 2.55 | 0.67 | 3.53 | 0.96 | 0.96
Results from laboratory 2.14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 906b (Certified Value 1.14 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [10 empty] | 2
| 1 | 1 | 0.61 | 1.26 | 1.3195 | 1.08 | 1.14 | 1.06 | 1 | 1.13 | 1.15 | 1.135 | 1.28 | 1.77 | 1.66 | 0.86 | 1.09 | < 0.1 | [10 empty] | 1
| 1 | 2 | 0.47 | 1.31 | 1.2888 | 1.07 | 1.15 | 1.08 | 1.09 | 1.07 | 1.12 | 1.1619 | 1.25 | 1.66 | 1.46 | 0.99 | 1.08 | < 0.1 | [10 empty] | 33
| 1 | 3 | 0.5 | 1.33 | 1.3848 | 1.12 | 1.11 | 1.08 | 1.05 | 1.04 | 1.12 | 1.1425 | 1.26 | 1.67 | 1.47 | 0.97 | 1.09 | < 0.1 | [10 empty] | 16
| 1 | 4 | 0.69 | 1.38 | 1.3073 | 1.08 | 1.14 | 1.06 | 1.06 | 1.02 | 1.12 | 1.0991 | 1.24 | 1.84 | 1.66 | 0.95 | 1.09 | < 0.1 | [10 empty] | 1.136
| 1 | 5 | 0.57 | 1.36 | 1.4158 | 1.1 | 1.13 | 1.06 | 1 | 1.09 | 1.11 | 1.1939 | 1.19 | 1.59 | 1.62 | 0.99 | 1.08 | < 0.1 | [10 empty] | 142
| 1 | 6 | 0.63 | 1.4 | 1.4153 | 1.01 | 1.18 | 1.14 | 1 | 1.07 | 1.13 | 1.1311 | 1.24 | 1.83 | 1.43 | 0.94 | 1.08 | < 0.1
| Mean |  | 0.57833 | 1.34 | 1.3552 | 1.0767 | 1.1417 | 1.08 | 1.0333 | 1.07 | 1.125 | 1.1439 | 1.2433 | 1.7267 | 1.55 | 0.95 | 1.085
| Median |  | 0.59 | 1.345 | 1.3522 | 1.08 | 1.14 | 1.07 | 1.025 | 1.07 | 1.12 | 1.1387 | 1.245 | 1.72 | 1.545 | 0.96 | 1.085
| Std Dev. |  | 0.082563 | 0.05099 | 0.056831 | 0.037238 | 0.023166 | 0.030984 | 0.038816 | 0.038471 | 0.013784 | 0.031878 | 0.030111 | 0.10172 | 0.1077 | 0.04858 | 0.0054772
| Rel.Std.Dev. |  | 14.28% | 3.805% | 4.193% | 3.459% | 2.029% | 2.869% | 3.756% | 3.595% | 1.225% | 2.787% | 2.422% | 5.891% | 6.949% | 5.114% | 0.5048%
| PDM3 |  | -49.09% | 17.96% | 19.3% | -5.224% | 0.4976% | -4.931% | -9.039% | -5.811% | -0.9695% | 0.6957% | 9.447% | 51.99% | 36.44% | -16.37% | -4.491%
| Z-Score (Absolute) |  | 3.38 | 1.51 | 1.61 | 0.18 | 0.24 | 0.16 | 0.46 | 0.22 | 0.13 | 0.25 | 0.89 | 3.99 | 2.86 | 1 | 0.13 | 6.78
| Analytical results for Hg in OREAS 906b (Indicative Value 0.015 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | [5 empty] | 3
| 1 | 1 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.03 | 0.01 | 0.02 | < 0.01 | 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | < 0.01 | 0.06 | < 0.01 | < 0.1 | 0.02 | 0.02 | 0.012 | [5 empty] | 1
| 1 | 2 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.01 | < 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.01 | 0.07 | < 0.01 | < 0.1 | 0.02 | 0.01 | 0.008 | [5 empty] | 14
| 1 | 3 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.01 | < 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.02 | 0.07 | < 0.01 | < 0.1 | 0.01 | 0.01 | 0.007 | [5 empty] | 16
| 1 | 4 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.02 | 0.01 | 0.02 | < 0.01 | < 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.02 | 0.06 | < 0.01 | < 0.1 | 0.02 | 0.03 | 0.005 | [5 empty] | 0.014806
| 1 | 5 | < 0.05 | 0.025 | < 5 | 0.01 | 0.02 | 0.01 | 0.02 | 0.01 | 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | < 0.01 | 0.07 | < 0.01 | < 0.1 | 0.02 | 0.03 | < 0.005 | [5 empty] | 26
| 1 | 6 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.03 | 0.01 | 0.02 | < 0.01 | < 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | < 0.01 | 0.07 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | < 0.005
| Mean |  |  | 0.025 |  | 0.01 | 0.02 | 0.01 | 0.02 | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.016667 | 0.066667 |  |  | 0.018 | 0.02 | 0.008
| Median |  |  | 0.025 |  | 0.01 | 0.02 | 0.01 | 0.02 | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.02 | 0.07 |  |  | 0.02 | 0.02 | 0.0075
| Std Dev. |  |  |  |  | 0 | 0.0089443 | 0 | 0 |  | 0 |  | 0 |  |  |  | 0.0057735 | 0.005164 |  |  | 0.0044721 | 0.01 | 0.0029439
| Rel.Std.Dev. |  |  |  |  | 0% | 44.72% | 0% | 0% |  | 0% |  | 0% |  |  |  | 34.64% | 7.746% |  |  | 24.85% | 50% | 36.8%
| PDM3 |  |  | 68.86% |  | -32.46% | 35.08% | -32.46% | 35.08% | -32.46% | -32.46% |  | -32.46% |  |  |  | 12.57% | 350.3% |  |  | 21.58% | 35.08% | -45.97%
| Z-Score (Absolute) |  | 0.45 | 0.67 | 148.8 | 0.45 | 0.15 | 0.45 | 0.15 | 0.7 | 0.65 | 28.92 | 0.45 | 28.92 | 28.92 | 58.89 | 0.4 | 2.95 | 0.75 | 1.95 | 0.1 | 0 | 0.68
| Analytical results for Ho in OREAS 906b (Indicative Value 0.23 ppm) | [27 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 0.221 | 0.19954 | 0.2798 | 0.3 | [22 empty] | 1
| 1 | 2 | 0.222 | 0.20425 | 0.25683 | 0.2 | [22 empty] | 16
| 1 | 3 | 0.222 | 0.20971 | 0.2662 | 0.3 | [22 empty] | 16
| 1 | 4 | 0.229 | 0.20195 | 0.2557 | 0.2 | [22 empty] | 0.23098
| 1 | 5 | 0.217 | 0.20537 | 0.2651 | 0.2 | [22 empty] | 22
| 1 | 6 | 0.23 | 0.20499 | 0.25301 | 0.2
| Mean |  | 0.2235 | 0.2043 | 0.26277 | 0.23333
| Median |  | 0.222 | 0.20462 | 0.26097 | 0.2
| Std Dev. |  | 0.00501 | 0.0034356 | 0.0098789 | 0.05164
| Rel.Std.Dev. |  | 2.242% | 1.682% | 3.759% | 22.13%
| PDM3 |  | -3.237% | -11.55% | 13.77% | 1.02%
| Z-Score (Absolute) |  | 0.23 | 1.12 | 1.6 | 0.23
| Analytical results for In in OREAS 906b (Certified Value 1.07 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [9 empty] | 3
| 1 | 1 | 1.05 | 1.09 | 1.029 | < 5 | 1.03 | 1.06 | 1.13 | 1.035 | 1.015 | 1.03 | 1.0521 | 0.93 | 1.07 | 1.112 | 1.031 | 1.06 | 1.1 | [9 empty] | 1
| 1 | 2 | 1.04 | 1.07 | 1.0691 | < 5 | 1.055 | 1.075 | 1.11 | 1.07 | 1.01 | 1.05 | 1.0748 | 0.97 | 1.09 | 1.044 | 1.098 | 1.08 | 1.11 | [9 empty] | 35
| 1 | 3 | 1.02 | 1.05 | 1.1071 | < 5 | 1.07 | 1.06 | 1.09 | 1.065 | 1.05 | 1.04 | 1.122 | 0.9 | 1.13 | 1.074 | 1.156 | 1.06 | 1.11 | [9 empty] | 16
| 1 | 4 | 1.07 | 1.07 | 1.0349 | < 5 | 1.025 | 1.06 | 1.115 | 1.045 | 0.999 | 1.03 | 1.0527 | 0.93 | 1.04 | 1.123 | 1.028 | 1.04 | 1.07 | [9 empty] | 1.0703
| 1 | 5 | 1.03 | 1.06 | 1.1107 | < 5 | 1.03 | 1.08 | 1.145 | 1.055 | 1.03 | 1.04 | 1.0865 | 0.93 | 1.12 | 1.114 | 1.09 | 1.08 | 1.09 | [9 empty] | 143
| 1 | 6 | 1.03 | 1.08 | 1.1511 | < 5 | 1.015 | 1.055 | 1.135 | 1.075 | 1.01 | 1.03 | 1.1049 | 0.97 | 1.14 | 1.112 | 1.068 | 1.06 | 1.09
| Mean |  | 1.04 | 1.07 | 1.0837 |  | 1.0375 | 1.065 | 1.1208 | 1.0575 | 1.019 | 1.0367 | 1.0822 | 0.93833 | 1.0983 | 1.0965 | 1.0785 | 1.0633 | 1.095
| Median |  | 1.035 | 1.07 | 1.0881 |  | 1.03 | 1.06 | 1.1225 | 1.06 | 1.0125 | 1.035 | 1.0807 | 0.93 | 1.105 | 1.112 | 1.079 | 1.06 | 1.095
| Std Dev. |  | 0.017889 | 0.014142 | 0.047747 |  | 0.020676 | 0.01 | 0.019854 | 0.015411 | 0.018221 | 0.008165 | 0.028099 | 0.027142 | 0.038687 | 0.030827 | 0.047824 | 0.015055 | 0.015166
| Rel.Std.Dev. |  | 1.72% | 1.322% | 4.406% |  | 1.993% | 0.939% | 1.771% | 1.457% | 1.788% | 0.7876% | 2.597% | 2.893% | 3.522% | 2.811% | 4.434% | 1.416% | 1.385%
| PDM3 |  | -2.831% | -0.02798% | 1.247% |  | -3.065% | -0.4951% | 4.721% | -1.196% | -4.793% | -3.142% | 1.11% | -12.33% | 2.619% | 2.448% | 0.7662% | -0.6509% | 2.308%
| Z-Score (Absolute) |  | 0.76 | 0 | 0.35 | 36.39 | 0.83 | 0.13 | 1.29 | 0.32 | 1.3 | 0.85 | 0.31 | 3.35 | 0.72 | 0.67 | 0.22 | 0.17 | 0.64
| Analytical results for K in OREAS 906b (Certified Value 0.228 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.2221 | 0.278 | 0.21236 | 0.21 | 0.95 | 0.21 | 0.21 | 0.25 | 0.21 | 0.23 | 0.216 | 0.214 | 0.21 | 0.2427 | 0.2 | 0.22 | 0.29254 | 0.21 | 0.462 | 0.226 | 0.23 | 0.29891 | 0.26 | 0.26 | 0.23 |  | 1
| 1 | 2 | 0.2263 | 0.275 | 0.21544 | 0.21 | 0.95 | 0.22 | 0.21 | 0.24 | 0.22 | 0.21 | 0.217 | 0.22 | 0.22 | 0.24187 | 0.21 | 0.21 | 0.29928 | 0.21 | 0.414 | 0.229 | 0.22 | 0.29784 | 0.26 | 0.27 | 0.22 |  | #N/A
| 1 | 3 | 0.2262 | 0.28 | 0.22377 | 0.25 | 0.96 | 0.22 | 0.21 | 0.24 | 0.21 | 0.21 | 0.225 | 0.21 | 0.22 | 0.25427 | 0.2 | 0.21 | 0.29752 | 0.19 | 0.411 | 0.236 | 0.23 | 0.28798 | 0.26 | 0.29 | 0.23 |  | 16
| 1 | 4 | 0.2296 | 0.283 | 0.21768 | 0.22 | 0.95 | 0.22 | 0.21 | 0.24 | 0.21 | 0.21 | 0.236 | 0.212 | 0.21 | 0.24897 | 0.21 | 0.21 | 0.29686 | 0.2 | 0.438 | 0.234 | 0.23 | 0.29744 | 0.26 | 0.27 | 0.23 |  | 0.22818
| 1 | 5 | 0.2196 | 0.28 | 0.22261 | 0.21 | 0.95 | 0.22 | 0.22 | 0.25 | 0.21 | 0.2 | 0.226 | 0.21 | 0.21 | 0.25399 | 0.21 | 0.22 | 0.29433 | 0.21 | 0.434 | 0.243 | 0.23 | 0.29869 | 0.26 | 0.26 | 0.22 |  | 144
| 1 | 6 | 0.2258 | 0.281 | 0.21148 | 0.27 | 0.95 | 0.21 | 0.21 | 0.25 | 0.21 | 0.21 | 0.232 | 0.22 | 0.22 | 0.24273 | 0.21 | 0.22 | 0.2983 | 0.19 | 0.454 | 0.23 | 0.23 | 0.29087 | 0.27 | 0.27 | 0.23
| Mean |  | 0.22493 | 0.2795 | 0.21722 | 0.22833 | 0.95167 | 0.21667 | 0.21167 | 0.245 | 0.21167 | 0.21167 | 0.22533 | 0.21433 | 0.215 | 0.24742 | 0.20667 | 0.215 | 0.29647 | 0.20167 | 0.4355 | 0.233 | 0.22833 | 0.29529 | 0.26167 | 0.27 | 0.22667
| Median |  | 0.226 | 0.28 | 0.21656 | 0.215 | 0.95 | 0.22 | 0.21 | 0.245 | 0.21 | 0.21 | 0.2255 | 0.213 | 0.215 | 0.24585 | 0.21 | 0.215 | 0.29719 | 0.205 | 0.436 | 0.232 | 0.23 | 0.29764 | 0.26 | 0.27 | 0.23
| Std Dev. |  | 0.0035348 | 0.0027386 | 0.0051368 | 0.025626 | 0.0040825 | 0.005164 | 0.0040825 | 0.0054772 | 0.0040825 | 0.0098319 | 0.0079415 | 0.0046332 | 0.0054772 | 0.0057921 | 0.005164 | 0.0054772 | 0.0025505 | 0.0098319 | 0.020569 | 0.0060663 | 0.0040825 | 0.0046639 | 0.0040825 | 0.010954 | 0.005164
| Rel.Std.Dev. |  | 1.571% | 0.9798% | 2.365% | 11.22% | 0.429% | 2.383% | 1.929% | 2.236% | 1.929% | 4.645% | 3.524% | 2.162% | 2.548% | 2.341% | 2.499% | 2.548% | 0.8603% | 4.875% | 4.723% | 2.604% | 1.788% | 1.579% | 1.56% | 4.057% | 2.278%
| PDM3 |  | -1.422% | 22.49% | -4.801% | 0.06794% | 317.1% | -5.045% | -7.236% | 7.372% | -7.236% | -7.236% | -1.247% | -6.068% | -5.775% | 8.433% | -9.428% | -5.775% | 29.93% | -11.62% | 90.86% | 2.113% | 0.06794% | 29.41% | 14.68% | 18.33% | -0.6625%
| Z-Score (Absolute) |  | 0.08 | 2.38 | 0.42 | 0.07 | 32.59 | 0.45 | 0.67 | 0.82 | 0.67 | 0.67 | 0.06 | 0.55 | 0.52 | 0.93 | 0.9 | 0.52 | 3.14 | 1.12 | 9.39 | 0.28 | 0.07 | 3.08 | 1.57 | 1.95 | 0
| Analytical results for La in OREAS 906b (Certified Value 37.6 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 2
| 1 | 1 | 39.133 | 38.7 | 36.0863 | 30.7 | 31.3 | 36.2 | 37.4 | 42.1 | 37 | 36 | 33.5 | 38 | 38.8 | 32.0719 | 34.3 | 34.8 | 40.1195 | 35 | 39.1 | 45 | 38.4 | 40 | 40.2 | 41.1 | 36.5 |  | 1
| 1 | 2 | 38.692 | 39.9 | 36.237 | 33.5 | 31.1 | 36.5 | 37.8 | 40.8 | 38.3 | 35.8 | 34.3 | 40 | 37.6 | 32.0738 | 34.6 | 34.1 | 40.1344 | 35 | 38.5 | 41.6 | 39.9 | 42 | 41.4 | 41.1 | 37.3 |  | 36
| 1 | 3 | 38.117 | 38.7 | 38.8429 | 34.4 | 31.7 | 38.4 | 37.4 | 41.8 | 37.9 | 36.6 | 34.3 | 40 | 39.2 | 32.5086 | 34 | 33.6 | 40.1068 | 35 | 39 | 43.3 | 39.9 | 42 | 40.9 | 41.2 | 36.8 |  | 16
| 1 | 4 | 39.795 | 39.9 | 37.4216 | 30.9 | 31.8 | 36.6 | 38.8 | 41.4 | 37.1 | 35.2 | 34.4 | 38 | 37.2 | 32.0465 | 35.3 | 35 | 40.1967 | 35 | 37 | 45.5 | 39.1 | 41 | 41.1 | 40.9 | 36.4 |  | 37.5884
| 1 | 5 | 37.093 | 39.8 | 39.1015 | 33 | 32.4 | 37.2 | 38.1 | 43.2 | 37.1 | 36.7 | 34.9 | 40 | 37.4 | 32.0235 | 34.3 | 34.3 | 40.0693 | 35 | 38.6 | 42 | 40.9 | 42 | 41 | 39.9 | 37.1 |  | 145
| 1 | 6 | 38.542 | 40 | 37.8938 | 33.9 | 31.6 | 36.8 | 37.3 | 42 | 38.2 | 35.3 | 34.7 | 40 | 38 | 31.7287 | 33.8 | 34.9 | 40.1695 | 35 | 39.8 | 43.2 | 39.9 | 42 | 40 | 41.3 | 36.5
| Mean |  | 38.562 | 39.5 | 37.5972 | 32.7333 | 31.65 | 36.95 | 37.8 | 41.8833 | 37.6 | 35.9333 | 34.35 | 39.3333 | 38.0333 | 32.0755 | 34.3833 | 34.45 | 40.1327 | 35 | 38.6667 | 43.4333 | 39.6833 | 41.5 | 40.7667 | 40.9167 | 36.7667
| Median |  | 38.617 | 39.85 | 37.6577 | 33.25 | 31.65 | 36.7 | 37.6 | 41.9 | 37.5 | 35.9 | 34.35 | 40 | 37.8 | 32.0592 | 34.3 | 34.55 | 40.1269 | 35 | 38.8 | 43.25 | 39.9 | 42 | 40.95 | 41.1 | 36.65
| Std Dev. |  | 0.91865 | 0.6229 | 1.2701 | 1.568 | 0.45056 | 0.78422 | 0.57619 | 0.80104 | 0.6 | 0.6314 | 0.48062 | 1.0328 | 0.80416 | 0.24958 | 0.52694 | 0.54681 | 0.045447 | 0 | 0.93737 | 1.5629 | 0.84951 | 0.83666 | 0.5465 | 0.51543 | 0.36697
| Rel.Std.Dev. |  | 2.382% | 1.577% | 3.378% | 4.79% | 1.424% | 2.122% | 1.524% | 1.913% | 1.596% | 1.757% | 1.399% | 2.626% | 2.114% | 0.7781% | 1.533% | 1.587% | 0.1132% | 0% | 2.424% | 3.598% | 2.141% | 2.016% | 1.341% | 1.26% | 0.9981%
| PDM3 |  | 2.59% | 5.086% | 0.02332% | -12.92% | -15.8% | -1.698% | 0.5629% | 11.43% | 0.03078% | -4.403% | -8.615% | 4.642% | 1.184% | -14.67% | -8.527% | -8.349% | 6.769% | -6.886% | 2.869% | 15.55% | 5.573% | 10.41% | 8.455% | 8.854% | -2.186%
| Z-Score (Absolute) |  | 0.22 | 0.49 | 0.06 | 1.46 | 1.78 | 0.25 | 0 | 1.18 | 0.06 | 0.54 | 1 | 0.44 | 0.07 | 1.65 | 0.99 | 0.97 | 0.67 | 0.81 | 0.25 | 1.63 | 0.54 | 1.07 | 0.86 | 0.9 | 0.3
| Analytical results for Li in OREAS 906b (Certified Value 3.27 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [5 empty] | 2
| 1 | 1 | 3.3 | 3.4 | 3.3308 | < 5 | 3.3 | 3.1 | 3.5 | 3 | 2.9 | 3 | 3 | < 10 | 2 | 2.87 | 14.48 | 3 | 5.7 | 5 | 4 | 3 | 3.8 | [5 empty] | 1
| 1 | 2 | 3.3 | 3.5 | 3.4171 | < 5 | 3.2 | 3 | 3.3 | 3.1 | 2.7 | 3 | 3 | < 10 | 2 | 2.93 | 14.41 | 3 | 6.1 | 5 | 4 | 3 | 4 | [5 empty] | #N/A
| 1 | 3 | 3.3 | 3.5 | 3.6505 | < 5 | 3.3 | 3 | 3.3 | 3.1 | 2.8 | 3 | 3 | < 10 | 2 | 2.8 | 14.64 | 3 | 9.9 | 5 | 4 | 3 | 4 | [5 empty] | 16
| 1 | 4 | 3.4 | 3.7 | 3.3601 | < 5 | 3.4 | 3.1 | 3.4 | 3.2 | 2.7 | 3 | 3 | < 10 | 2 | 2.97 | 14.42 | 3 | 5.9 | 5 | 4 | 3 | 3.8 | [5 empty] | 3.267
| 1 | 5 | 3.2 | 3.6 | 3.6257 | < 5 | 3.4 | 3.1 | 3.4 | 3 | 2.7 | 3 | 3 | < 10 | 2 | 3.13 | 14.51 | 3 | 5.7 | 5 | 4 | 3 | 3.8 | [5 empty] | 146
| 1 | 6 | 3.4 | 3.6 | 3.407 | < 5 | 3.1 | 2.9 | 3.3 | 3.1 | 2.7 | 3 | 3 | < 10 | 3 | 2.93 | 14.24 | 3 | 7.9 | 5 | 4 | 3 | 3.9
| Mean |  | 3.3167 | 3.55 | 3.4652 |  | 3.2833 | 3.0333 | 3.3667 | 3.0833 | 2.75 | 3 | 3 |  | 2.1667 | 2.9383 | 14.45 | 3 | 6.8667 | 5 | 4 | 3 | 3.8833
| Median |  | 3.3 | 3.55 | 3.412 |  | 3.3 | 3.05 | 3.35 | 3.1 | 2.7 | 3 | 3 |  | 2 | 2.93 | 14.45 | 3 | 6 | 5 | 4 | 3 | 3.85
| Std Dev. |  | 0.075277 | 0.10488 | 0.13779 |  | 0.1169 | 0.08165 | 0.08165 | 0.075277 | 0.083666 | 0 | 0 |  | 0.40825 | 0.11107 | 0.13206 | 0 | 1.7037 | 0 | 0 | 0 | 0.098319
| Rel.Std.Dev. |  | 2.27% | 2.954% | 3.976% |  | 3.561% | 2.692% | 2.425% | 2.441% | 3.042% | 0% | 0% |  | 18.84% | 3.78% | 0.9139% | 0% | 24.81% | 0% | 0% | 0% | 2.532%
| PDM3 |  | 1.52% | 8.662% | 6.066% |  | 0.4993% | -7.153% | 3.05% | -5.622% | -15.83% | -8.173% | -8.173% |  | -33.68% | -10.06% | 342.3% | -8.173% | 110.2% | 53.04% | 22.44% | -8.173% | 18.86%
| Z-Score (Absolute) |  | 0.05 | 0.39 | 0.23 | 1.59 | 0.11 | 0.58 | 0.05 | 0.49 | 1.12 | NA | NA | 3.14 | NA | 0.76 | 21.05 | NA | 6.68 | NA | NA | NA | 1.03
Results from laboratories 2.12, 2.14, 2.16, 2.19, 2.21, 2.22 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 906b (Indicative Value 0.031 ppm) | [27 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 2.01 | 2.06 | 2.14 | 2.15 | 2.21 | 2.24 | 2.25 | [19 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [19 empty] | 3
|  |  | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.5g | [19 empty] | 3
| 1 | 1 | 0.027 | < 5 | 0.03 | < 0.05 | 0.04 | 0.03 | < 0.1 | [19 empty] | 1
| 1 | 2 | 0.028 | < 5 | 0.03 | < 0.05 | 0.04 | 0.03 | < 0.1 | [19 empty] | 9
| 1 | 3 | 0.027 | < 5 | 0.03 | < 0.05 | 0.04 | 0.03 | < 0.1 | [19 empty] | 16
| 1 | 4 | 0.028 | < 5 | 0.02 | < 0.05 | 0.04 | 0.03 | < 0.1 | [19 empty] | 0.031333
| 1 | 5 | 0.025 | < 5 | 0.03 | < 0.05 | 0.05 | 0.03 | < 0.1 | [19 empty] | 23
| 1 | 6 | 0.027 | < 5 | 0.02 | < 0.05 | 0.04 | 0.03 | < 0.1
| Mean |  | 0.027 |  | 0.026667 |  | 0.041667 | 0.03
| Median |  | 0.027 |  | 0.03 |  | 0.04 | 0.03
| Std Dev. |  | 0.0010954 |  | 0.005164 |  | 0.0040825 | 0
| Rel.Std.Dev. |  | 4.057% |  | 19.36% |  | 9.798% | 0%
| PDM3 |  | -13.83% |  | -14.89% |  | 32.98% | -4.255%
| Z-Score (Absolute) |  | 0.4 | 333.11 | 0.45 | 0.67 | 1.57 | 0 | 2.7
| Analytical results for Mg in OREAS 906b (Certified Value 0.097 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.1 | 0.1 | 0.096327 | 0.09 | 0.19 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.096 | 0.096 | 0.1 | 0.097492 | 0.105 | 0.1 | 0.096295 | 0.09 | 0.115 | 0.091 | 0.11 | 0.1094 | 0.1 | 0.1 | 0.1 |  | 1
| 1 | 2 | 0.1 | 0.1 | 0.097226 | 0.09 | 0.2 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.098 | 0.09 | 0.09 | 0.095595 | 0.106 | 0.09 | 0.096695 | 0.09 | 0.106 | 0.093 | 0.1 | 0.1089 | 0.1 | 0.1 | 0.1 |  | #N/A
| 1 | 3 | 0.1 | 0.1 | 0.10334 | 0.1 | 0.19 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.098 | 0.09 | 0.1 | 0.0965 | 0.105 | 0.09 | 0.096081 | 0.09 | 0.108 | 0.094 | 0.1 | 0.1064 | 0.1 | 0.1 | 0.1 |  | 16
| 1 | 4 | 0.1 | 0.1 | 0.096877 | 0.09 | 0.19 | 0.1 | 0.09 | 0.1 | 0.09 | 0.09 | 0.1 | 0.094 | 0.1 | 0.096916 | 0.107 | 0.1 | 0.096845 | 0.09 | 0.106 | 0.092 | 0.11 | 0.1147 | 0.1 | 0.1 | 0.1 |  | 0.097014
| 1 | 5 | 0.1 | 0.1 | 0.10038 | 0.09 | 0.19 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.096 | 0.1 | 0.09 | 0.097439 | 0.106 | 0.09 | 0.096763 | 0.09 | 0.122 | 0.097 | 0.11 | 0.1072 | 0.11 | 0.1 | 0.1 |  | 147
| 1 | 6 | 0.1 | 0.1 | 0.10155 | 0.09 | 0.19 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.099 | 0.1 | 0.1 | 0.096774 | 0.107 | 0.09 | 0.096101 | 0.09 | 0.11 | 0.094 | 0.1 | 0.1098 | 0.1 | 0.1 | 0.1
| Mean |  | 0.1 | 0.1 | 0.099283 | 0.091667 | 0.19167 | 0.091667 | 0.09 | 0.1 | 0.09 | 0.09 | 0.097833 | 0.095 | 0.096667 | 0.096786 | 0.106 | 0.093333 | 0.096463 | 0.09 | 0.11117 | 0.0935 | 0.105 | 0.1094 | 0.10167 | 0.1 | 0.1
| Median |  | 0.1 | 0.1 | 0.098803 | 0.09 | 0.19 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.098 | 0.095 | 0.1 | 0.096845 | 0.106 | 0.09 | 0.096495 | 0.09 | 0.109 | 0.0935 | 0.105 | 0.10915 | 0.1 | 0.1 | 0.1
| Std Dev. |  | 1.5202e-17 | 1.5202e-17 | 0.0028823 | 0.0040825 | 0.0040825 | 0.0040825 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 0.0016021 | 0.0045166 | 0.005164 | 0.00069867 | 0.00089443 | 0.005164 | 0.00034495 | 1.5202e-17 | 0.0062743 | 0.0020736 | 0.0054772 | 0.00291 | 0.0040825 | 1.5202e-17 | 1.5202e-17
| Rel.Std.Dev. |  | 1.52e-14% | 1.52e-14% | 2.903% | 4.454% | 2.13% | 4.454% | 1.689e-14% | 1.52e-14% | 1.689e-14% | 1.689e-14% | 1.638% | 4.754% | 5.342% | 0.7219% | 0.8438% | 5.533% | 0.3576% | 1.689e-14% | 5.644% | 2.218% | 5.216% | 2.66% | 4.016% | 1.52e-14% | 1.52e-14%
| PDM3 |  | 3.078% | 3.078% | 2.339% | -5.512% | 97.57% | -5.512% | -7.23% | 3.078% | -7.23% | -7.23% | 0.8446% | -2.076% | -0.358% | -0.235% | 9.263% | -3.794% | -0.5676% | -7.23% | 14.59% | -3.622% | 8.232% | 12.77% | 4.796% | 3.078% | 3.078%
| Z-Score (Absolute) |  | NA | NA | 0.37 | 0.66 | 12.83 | 0.66 | 0.89 | NA | 0.89 | 0.89 | 0.17 | 0.21 | 0.01 | 0.03 | 1.27 | 0.44 | 0.01 | 0.89 | 1.97 | 0.41 | 1.14 | 1.73 | 0.69 | NA | NA
Results from laboratories 2.01, 2.02, 2.09, 2.25 and 2.27 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Mn in OREAS 906b (Certified Value 0.028 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.0278 | 0.0301 | 0.027802 | 0.0275 | 0.03174 | 0.0279 | 0.0279 | 0.03 | 0.0278 | 0.0277 | 0.0264 | 0.0291 | 0.028 | 0.028212 | 0.0262 | 0.0276 | 0.027809 | 0.0278 | 0.0244 | 0.0277 | 0.0287 | 0.0348 | 0.0311 | 0.0287 | 0.0296 |  | 1
| 1 | 2 | 0.0287 | 0.0289 | 0.027185 | 0.0282 | 0.03329 | 0.0285 | 0.0278 | 0.0299 | 0.0285 | 0.0278 | 0.0277 | 0.0331 | 0.0284 | 0.028292 | 0.0258 | 0.0267 | 0.027125 | 0.028 | 0.0224 | 0.0291 | 0.0287 | 0.03588 | 0.0315 | 0.0293 | 0.0293 |  | 38
| 1 | 3 | 0.0282 | 0.0293 | 0.02737 | 0.0305 | 0.03169 | 0.029 | 0.0272 | 0.0293 | 0.0285 | 0.0279 | 0.0268 | 0.0294 | 0.0285 | 0.028334 | 0.0262 | 0.0264 | 0.027862 | 0.0281 | 0.0234 | 0.0294 | 0.0283 | 0.03575 | 0.0311 | 0.0294 | 0.0293 |  | 16
| 1 | 4 | 0.0284 | 0.029 | 0.02725 | 0.0284 | 0.03201 | 0.0284 | 0.0271 | 0.03 | 0.0291 | 0.0277 | 0.0279 | 0.0297 | 0.0283 | 0.028434 | 0.0259 | 0.0276 | 0.027836 | 0.0278 | 0.0229 | 0.0279 | 0.0288 | 0.03406 | 0.031 | 0.0279 | 0.0294 |  | 0.028375
| 1 | 5 | 0.0279 | 0.0289 | 0.027774 | 0.0286 | 0.03259 | 0.0288 | 0.0279 | 0.0301 | 0.0278 | 0.0275 | 0.026 | 0.0296 | 0.0288 | 0.028938 | 0.0261 | 0.0265 | 0.027334 | 0.0279 | 0.0268 | 0.0297 | 0.0288 | 0.03523 | 0.0311 | 0.0286 | 0.0291 |  | 148
| 1 | 6 | 0.0278 | 0.0289 | 0.027619 | 0.0296 | 0.03232 | 0.0279 | 0.0268 | 0.0303 | 0.0287 | 0.0279 | 0.0269 | 0.0299 | 0.0284 | 0.02815 | 0.0261 | 0.0276 | 0.027794 | 0.028 | 0.0229 | 0.0296 | 0.03 | 0.03497 | 0.0312 | 0.0294 | 0.0291
| Mean |  | 0.028133 | 0.029183 | 0.0275 | 0.0288 | 0.032273 | 0.028417 | 0.02745 | 0.029933 | 0.0284 | 0.02775 | 0.02695 | 0.030133 | 0.0284 | 0.028393 | 0.02605 | 0.027067 | 0.027627 | 0.027933 | 0.0238 | 0.0289 | 0.028883 | 0.035115 | 0.031167 | 0.028883 | 0.0293
| Median |  | 0.02805 | 0.02895 | 0.027494 | 0.0285 | 0.032165 | 0.02845 | 0.0275 | 0.03 | 0.0285 | 0.02775 | 0.02685 | 0.02965 | 0.0284 | 0.028313 | 0.0261 | 0.02715 | 0.027801 | 0.02795 | 0.02315 | 0.02925 | 0.02875 | 0.0351 | 0.0311 | 0.029 | 0.0293
| Std Dev. |  | 0.00036697 | 0.00047504 | 0.00026782 | 0.0010752 | 0.000605 | 0.0004535 | 0.00047645 | 0.00033862 | 0.00051381 | 0.00015166 | 0.00073417 | 0.0014787 | 0.00026077 | 0.00028416 | 0.00016432 | 0.00059217 | 0.00031564 | 0.00012111 | 0.0016186 | 0.00087864 | 0.00057764 | 0.00066857 | 0.00017512 | 0.00059805 | 0.00018974
| Rel.Std.Dev. |  | 1.304% | 1.628% | 0.9739% | 3.733% | 1.875% | 1.596% | 1.736% | 1.131% | 1.809% | 0.5465% | 2.724% | 4.907% | 0.9182% | 1.001% | 0.6308% | 2.188% | 1.143% | 0.4336% | 6.801% | 3.04% | 2% | 1.904% | 0.5619% | 2.071% | 0.6476%
| PDM3 |  | -0.8522% | 2.848% | -3.085% | 1.497% | 13.74% | 0.1464% | -3.26% | 5.491% | 0.08762% | -2.203% | -5.022% | 6.196% | 0.08762% | 0.06385% | -8.194% | -4.611% | -2.638% | -1.557% | -16.12% | 1.85% | 1.791% | 23.75% | 9.838% | 1.791% | 3.259%
| Z-Score (Absolute) |  | 0.23 | 0.67 | 0.77 | 0.34 | 3.33 | 0.01 | 0.82 | 1.32 | 0 | 0.56 | 1.25 | 1.49 | 0 | 0.01 | 2.02 | 1.15 | 0.67 | 0.4 | 3.96 | 0.43 | 0.42 | 5.78 | 2.38 | 0.42 | 0.77
| Analytical results for Mo in OREAS 906b (Certified Value 4.47 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 4.7 | 4.5 | 4.1013 | 3.72 | 3.6 | 4.36 | 4.7 | 4.44 | 4.66 | 4.4 | 4.32 | 5.2 | 4.33 | 4.8397 | 4 | 4.67 | 5 | 3.9 | 6.12 | 4.13 | 3.92 | 4.6 | 4.67 | 4 | 4.47 |  | 1
| 1 | 2 | 4.9 | 4.5 | 4.6608 | 4.58 | 3.6 | 4.56 | 4.66 | 4.45 | 4.79 | 4.56 | 4.49 | 5 | 4.54 | 4.7637 | 4 | 4.84 | 5 | 3.9 | 5.42 | 4.12 | 4.82 | 4.59 | 4.6 | 3.9 | 4.33 |  | 39
| 1 | 3 | 4.8 | 4.4 | 4.7077 | 4.24 | 3.3 | 4.63 | 4.52 | 4.43 | 4.69 | 4.66 | 4.61 | 5 | 4.37 | 4.6994 | 4 | 4.94 | 5 | 4 | 5.28 | 4.38 | 3.93 | 4.55 | 4.55 | 3.8 | 4.41 |  | 16
| 1 | 4 | 4.9 | 4.5 | 4.5255 | 3.81 | 3.5 | 7.43 | 4.37 | 4.4 | 4.62 | 4.5 | 4.5 | 5.1 | 4.33 | 4.9614 | 5 | 4.76 | 5 | 3.6 | 6.11 | 4.24 | 3.8 | 4.54 | 4.61 | 3.9 | 4.44 |  | 4.4654
| 1 | 5 | 4.8 | 4.6 | 4.4635 | 4.06 | 3.6 | 4.47 | 4.67 | 4.61 | 4.76 | 4.68 | 4.53 | 5 | 4.43 | 4.7361 | 4 | 4.76 | 5 | 3.8 | 5.49 | 4.36 | 3.7 | 4.79 | 4.66 | 4 | 4.44 |  | 149
| 1 | 6 | 4.8 | 4.6 | 4.1833 | 4.07 | 3.6 | 4.68 | 4.52 | 4.55 | 4.77 | 4.34 | 4.42 | 5 | 4.36 | 4.5073 | 4 | 4.82 | 5 | 3.8 | 5.81 | 4.08 | 3.8 | 4.63 | 4.65 | 3.9 | 4.39
| Mean |  | 4.8167 | 4.5167 | 4.4403 | 4.08 | 3.5333 | 5.0217 | 4.5733 | 4.48 | 4.715 | 4.5233 | 4.4783 | 5.05 | 4.3933 | 4.7512 | 4.1667 | 4.7983 | 5 | 3.8333 | 5.705 | 4.2183 | 3.995 | 4.6167 | 4.6233 | 3.9167 | 4.4133
| Median |  | 4.8 | 4.5 | 4.4945 | 4.065 | 3.6 | 4.595 | 4.59 | 4.445 | 4.725 | 4.53 | 4.495 | 5 | 4.365 | 4.7499 | 4 | 4.79 | 5 | 3.85 | 5.65 | 4.185 | 3.86 | 4.595 | 4.63 | 3.9 | 4.425
| Std Dev. |  | 0.075277 | 0.075277 | 0.24855 | 0.30939 | 0.12111 | 1.1854 | 0.12644 | 0.081486 | 0.068337 | 0.13706 | 0.099079 | 0.083666 | 0.080664 | 0.15137 | 0.40825 | 0.091305 | 0 | 0.13663 | 0.36204 | 0.12906 | 0.41317 | 0.091141 | 0.045461 | 0.075277 | 0.049261
| Rel.Std.Dev. |  | 1.563% | 1.667% | 5.598% | 7.583% | 3.428% | 23.61% | 2.765% | 1.819% | 1.449% | 3.03% | 2.212% | 1.657% | 1.836% | 3.186% | 9.798% | 1.903% | 0% | 3.564% | 6.346% | 3.06% | 10.34% | 1.974% | 0.9833% | 1.922% | 1.116%
| PDM3 |  | 7.867% | 1.148% | -0.5611% | -8.631% | -20.87% | 12.46% | 2.417% | 0.3271% | 5.59% | 1.298% | 0.2898% | 13.09% | -1.614% | 6.402% | -6.69% | 7.456% | 11.97% | -14.15% | 27.76% | -5.533% | -10.53% | 3.388% | 3.537% | -12.29% | -1.166%
| Z-Score (Absolute) |  | 0.86 | 0 | 0.22 | 1.26 | 2.83 | 1.45 | 0.16 | 0.11 | 0.57 | 0.02 | 0.11 | 1.53 | 0.35 | 0.67 | NA | 0.81 | NA | 1.96 | 3.42 | 0.86 | 1.5 | 0.29 | 0.31 | 1.72 | 0.3
Results from laboratories 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 906b (Certified Value 0.07 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.06 | 0.07 | 0.074074 | 0.067 | 0.08 | 0.06 | 0.06 | 0.08 | 0.09 | 0.08 | 0.07 | 0.157 | 0.06 | 0.066436 | 0.062 | 0.07 | 0.07663 | 0.07 | 0.0699 | 0.086 | 0.07 | 0.0973 | 0.07 | 0.061 | 0.069 |  | 1
| 1 | 2 | 0.06 | 0.08 | 0.066501 | 0.069 | 0.07 | 0.06 | 0.06 | 0.08 | 0.09 | 0.06 | 0.07 | 0.16 | 0.06 | 0.067424 | 0.063 | 0.07 | 0.076455 | 0.07 | 0.0623 | 0.078 | 0.06 | 0.0995 | 0.07 | 0.065 | 0.067 |  | 10
| 1 | 3 | 0.06 | 0.08 | 0.073316 | 0.095 | 0.08 | 0.07 | 0.06 | 0.08 | 0.09 | 0.07 | 0.08 | 0.16 | 0.06 | 0.06999 | 0.063 | 0.07 | 0.076027 | 0.07 | 0.0658 | 0.081 | 0.07 | 0.0922 | 0.07 | 0.064 | 0.068 |  | 16
| 1 | 4 | 0.06 | 0.08 | 0.068228 | 0.069 | 0.07 | 0.07 | 0.07 | 0.08 | 0.09 | 0.07 | 0.08 | 0.155 | 0.07 | 0.068249 | 0.065 | 0.07 | 0.075976 | 0.07 | 0.0675 | 0.091 | 0.07 | 0.095 | 0.07 | 0.066 | 0.068 |  | 0.069706
| 1 | 5 | 0.06 | 0.08 | 0.075848 | 0.068 | 0.07 | 0.07 | 0.06 | 0.08 | 0.09 | 0.06 | 0.07 | 0.16 | 0.07 | 0.069598 | 0.062 | 0.07 | 0.076345 | 0.07 | 0.0759 | 0.088 | 0.07 | 0.0979 | 0.07 | 0.063 | 0.066 |  | 150
| 1 | 6 | 0.06 | 0.08 | 0.068916 | 0.101 | 0.08 | 0.06 | 0.05 | 0.08 | 0.09 | 0.07 | 0.07 | 0.16 | 0.06 | 0.068086 | 0.063 | 0.07 | 0.076129 | 0.07 | 0.0713 | 0.1 | 0.07 | 0.0969 | 0.07 | 0.064 | 0.067
| Mean |  | 0.06 | 0.078333 | 0.071147 | 0.078167 | 0.075 | 0.065 | 0.06 | 0.08 | 0.09 | 0.068333 | 0.073333 | 0.15867 | 0.063333 | 0.068297 | 0.063 | 0.07 | 0.07626 | 0.07 | 0.068783 | 0.087333 | 0.068333 | 0.096467 | 0.07 | 0.063833 | 0.0675
| Median |  | 0.06 | 0.08 | 0.071116 | 0.069 | 0.075 | 0.065 | 0.06 | 0.08 | 0.09 | 0.07 | 0.07 | 0.16 | 0.06 | 0.068168 | 0.063 | 0.07 | 0.076237 | 0.07 | 0.0687 | 0.087 | 0.07 | 0.0971 | 0.07 | 0.064 | 0.0675
| Std Dev. |  | 0 | 0.0040825 | 0.0037538 | 0.015497 | 0.0054772 | 0.0054772 | 0.0063246 | 0 | 1.5202e-17 | 0.0075277 | 0.005164 | 0.0021602 | 0.005164 | 0.001329 | 0.0010954 | 0 | 0.00025841 | 0 | 0.0047076 | 0.0077889 | 0.0040825 | 0.0025492 | 0 | 0.0017224 | 0.0010488
| Rel.Std.Dev. |  | 0% | 5.212% | 5.276% | 19.83% | 7.303% | 8.427% | 10.54% | 0% | 1.689e-14% | 11.02% | 7.042% | 1.362% | 8.154% | 1.946% | 1.739% | 0% | 0.3389% | 0% | 6.844% | 8.919% | 5.974% | 2.643% | 0% | 2.698% | 1.554%
| PDM3 |  | -13.92% | 12.38% | 2.067% | 12.14% | 7.594% | -6.752% | -13.92% | 14.77% | 29.11% | -1.97% | 5.203% | 127.6% | -9.143% | -2.021% | -9.621% | 0.4214% | 9.402% | 0.4214% | -1.324% | 25.29% | -1.97% | 38.39% | 0.4214% | -8.425% | -3.165%
| Z-Score (Absolute) |  | 1.09 | 0.91 | 0.13 | 0.89 | 0.55 | 0.55 | 1.09 | 1.09 | 2.19 | 0.18 | 0.36 | 9.7 | 0.73 | 0.19 | 0.77 | 0 | 0.68 | 0 | 0.13 | 1.9 | 0.18 | 2.89 | 0 | 0.67 | 0.27
| Analytical results for Nb in OREAS 906b (Certified Value 0.47 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.22 | 0.6 | 0.54315 | 10.9 | 0.35 | 0.33 | 0.27 | 0.29 | 0.31 | 0.63 | 0.76816 | < 1 | 0.73 | 0.63 | 0.54 | 0.48 | 0.2 | [9 empty] | 1
| 1 | 2 | 0.2 | 0.6 | 0.55273 | 10.9 | 0.31 | 0.35 | 0.27 | 0.36 | 0.32 | 0.67 | 0.72832 | < 1 | 0.7 | 0.61 | 0.55 | 0.48 | 0.3 | [9 empty] | 40
| 1 | 3 | 0.21 | 0.6 | 0.58541 | 10.9 | 0.33 | 0.35 | 0.29 | 0.41 | 0.3 | 0.63 | 0.70308 | < 1 | 0.67 | 0.64 | 0.52 | 0.46 | 0.3 | [9 empty] | 16
| 1 | 4 | 0.25 | 0.6 | 0.55505 | 10.9 | 0.32 | 0.32 | 0.24 | 0.31 | 0.3 | 0.66 | 0.70642 | < 1 | 0.76 | 0.67 | 0.56 | 0.49 | 0.3 | [9 empty] | 0.46902
| 1 | 5 | 0.22 | 0.6 | 0.58293 | 10.9 | 0.32 | 0.34 | 0.28 | 0.29 | 0.34 | 0.66 | 0.7074 | < 1 | 0.76 | 0.65 | 0.54 | 0.45 | 0.3 | [9 empty] | 151
| 1 | 6 | 0.23 | 0.6 | 0.54494 | 10.9 | 0.3 | 0.35 | 0.27 | 0.28 | 0.32 | 0.65 | 0.70401 | < 1 | 0.75 | 0.6 | 0.59 | 0.48 | 0.3
| Mean |  | 0.22167 | 0.6 | 0.5607 | 10.9 | 0.32167 | 0.34 | 0.27 | 0.32333 | 0.315 | 0.65 | 0.71956 |  | 0.72833 | 0.63333 | 0.55 | 0.47333 | 0.28333
| Median |  | 0.22 | 0.6 | 0.55389 | 10.9 | 0.32 | 0.345 | 0.27 | 0.3 | 0.315 | 0.655 | 0.70691 |  | 0.74 | 0.635 | 0.545 | 0.48 | 0.3
| Std Dev. |  | 0.017224 | 0 | 0.018744 | 0 | 0.017224 | 0.012649 | 0.016733 | 0.051251 | 0.015166 | 0.016733 | 0.025582 |  | 0.03656 | 0.02582 | 0.023664 | 0.015055 | 0.040825
| Rel.Std.Dev. |  | 7.77% | 0% | 3.343% | 0% | 5.355% | 3.72% | 6.197% | 15.85% | 4.815% | 2.574% | 3.555% |  | 5.02% | 4.077% | 4.303% | 3.181% | 14.41%
| PDM3 |  | -52.74% | 27.93% | 19.55% | 2224% | -31.42% | -27.51% | -42.43% | -31.06% | -32.84% | 38.59% | 53.42% |  | 55.29% | 35.03% | 17.27% | 0.9204% | -39.59%
| Z-Score (Absolute) |  | 1.06 | NA | 0.23 | 39.7 | 0.68 | 0.61 | 0.88 | 0.67 | 0.71 | 0.57 | 0.84 | 0 | 0.87 | 0.51 | 0.19 | 0.1 | NA
Results from laboratories 2.02 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 906b (Indicative Value 32 ppm) | [27 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 1
| 1 | 1 | 32.069 | 29.6565 | 34.6153 | 32.2 | [22 empty] | 1
| 1 | 2 | 31.902 | 29.5805 | 34.516 | 32.4 | [22 empty] | 18
| 1 | 3 | 31.482 | 29.7479 | 34.3195 | 32.5 | [22 empty] | 16
| 1 | 4 | 32.456 | 29.2911 | 34.4489 | 32.2 | [22 empty] | 32.0002
| 1 | 5 | 30.683 | 29.6415 | 34.7079 | 31.3 | [22 empty] | 24
| 1 | 6 | 31.638 | 29.3261 | 34.8226 | 32.5
| Mean |  | 31.705 | 29.5406 | 34.5717 | 32.1833
| Median |  | 31.77 | 29.611 | 34.5657 | 32.3
| Std Dev. |  | 0.60564 | 0.18785 | 0.18176 | 0.4535
| Rel.Std.Dev. |  | 1.91% | 0.6359% | 0.5258% | 1.409%
| PDM3 |  | -0.9224% | -7.686% | 8.036% | 0.5724%
| Z-Score (Absolute) |  | 0.12 | 1.23 | 1.34 | 0.12
| Analytical results for Ni in OREAS 906b (Certified Value 2.69 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 3 | 3 | 2.6827 | 2.5 | 1.95 | 2.5 | 2.8 | 2.3 | 2.7 | 2.6 | 2.5 | 2.7 | 2.7 | 2.6996 | 3 | 2.7 | 3 | 2.5 | 2.6 | 4.5 | 2.9 | 2.8 | 2.8 |  |  |  | 1
| 1 | 2 | 3.1 | 3 | 2.3522 | 3 | 2.13 | 3 | 2.8 | 2.4 | 2.9 | 2.6 | 2.5 | 3 | 2.7 | 3.3405 | 3 | 2.9 | 3 | 2.7 | 2.7 | 4.2 | 2.7 | 2.5 | 2.8 |  |  |  | 11
| 1 | 3 | 2.7 | 3 | 2.7436 | 2.8 | 2.16 | 2.7 | 2.7 | 2.5 | 3.2 | 2.5 | 2.5 | 3 | 2.8 | 3.2488 | 3 | 2.5 | 3 | 2.7 | 2.6 | 4.2 | 3 | 2.5 | 2.8 |  |  |  | 16
| 1 | 4 | 2.8 | 3 | 2.86 | 2.5 | 1.93 | 2.6 | 2.7 | 2.7 | 2.6 | 2.6 | 2.5 | 3.7 | 2.6 | 3.4219 | 3 | 2.7 | 2 | 2.4 | 2.6 | 4.4 | 2.8 | 2.3 | 2.7 |  |  |  | 2.6921
| 1 | 5 | 3.2 | 3 | 2.3767 | 2.7 | 1.91 | 2.5 | 2.9 | 2.5 | 2.6 | 2.7 | 2.3 | 3 | 2.8 | 2.9742 | 3 | 2.8 | 2 | 2.5 | 2.6 | 4.3 | 2.9 | 2.7 | 2.7 |  |  |  | 152
| 1 | 6 | 2.7 | 3 | 2.7796 | 2.7 | 2.07 | 3.6 | 2.7 | 2.5 | 2.8 | 2.7 | 2.3 | 3 | 2.6 | 3.1742 | 2 | 2.6 | 3 | 2.4 | 2.7 | 4.6 | 3.1 | 2.8 | 2.7
| Mean |  | 2.9167 | 3 | 2.6325 | 2.7 | 2.025 | 2.8167 | 2.7667 | 2.4833 | 2.8 | 2.6167 | 2.4333 | 3.0667 | 2.7 | 3.1432 | 2.8333 | 2.7 | 2.6667 | 2.5333 | 2.6333 | 4.3667 | 2.9 | 2.6 | 2.75
| Median |  | 2.9 | 3 | 2.7131 | 2.7 | 2.01 | 2.65 | 2.75 | 2.5 | 2.75 | 2.6 | 2.5 | 3 | 2.7 | 3.2115 | 3 | 2.7 | 3 | 2.5 | 2.6 | 4.35 | 2.9 | 2.6 | 2.75
| Std Dev. |  | 0.2137 | 0 | 0.21556 | 0.18974 | 0.10877 | 0.42622 | 0.08165 | 0.13292 | 0.22804 | 0.075277 | 0.10328 | 0.33267 | 0.089443 | 0.26613 | 0.40825 | 0.14142 | 0.5164 | 0.13663 | 0.05164 | 0.1633 | 0.14142 | 0.2 | 0.054772
| Rel.Std.Dev. |  | 7.327% | 0% | 8.189% | 7.027% | 5.371% | 15.13% | 2.951% | 5.352% | 8.144% | 2.877% | 4.244% | 10.85% | 3.313% | 8.467% | 14.41% | 5.238% | 19.36% | 5.393% | 1.961% | 3.74% | 4.877% | 7.692% | 1.992%
| PDM3 |  | 8.341% | 11.44% | -2.215% | 0.2932% | -24.78% | 4.627% | 2.77% | -7.755% | 4.008% | -2.802% | -9.612% | 13.91% | 0.2932% | 16.76% | 5.246% | 0.2932% | -0.945% | -5.898% | -2.183% | 62.2% | 7.722% | -3.421% | 2.151%
| Z-Score (Absolute) |  | 1.35 | NA | 0.42 | 0 | 4.2 | 0.73 | 0.41 | 1.35 | 0.62 | 0.52 | 1.66 | 2.28 | 0 | 2.76 | NA | 0 | NA | 1.04 | 0.41 | 10.37 | 1.24 | 0.62 | 0.31
Results from laboratories 2.02, 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 906b (Certified Value 0.025 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.0236 | 0.0268 | 0.022804 | 0.023 | 0.02 | 0.025 | 0.026 | 0.027 | 0.024 | 0.024 | 0.024 | 0.0277 | 0.025 | 0.025538 | 0.023 | 0.024967 | 0.0247 | 0.0282 | 0.021 | 0.0237 | 0.0248 | 0.03 | 0.023 | 0.023 |  |  | 1
| 1 | 2 | 0.0238 | 0.0266 | 0.023323 | 0.024 | 0.02 | 0.025 | 0.026 | 0.026 | 0.025 | 0.024 | 0.025 | 0.029 | 0.0241 | 0.02533 | 0.023 | 0.0245 | 0.0246 | 0.0221 | 0.021 | 0.0244 | 0.0236 | 0.03 | 0.023 | 0.023 |  |  | #N/A
| 1 | 3 | 0.024 | 0.0269 | 0.024445 | 0.025 | 0.02 | 0.026 | 0.026 | 0.026 | 0.024 | 0.024 | 0.025 | 0.028 | 0.0247 | 0.025364 | 0.023 | 0.024433 | 0.0244 | 0.0245 | 0.022 | 0.0244 | 0.0241 | 0.03 | 0.023 | 0.023 |  |  | 16
| 1 | 4 | 0.0238 | 0.0272 | 0.023306 | 0.024 | 0.02 | 0.025 | 0.024 | 0.027 | 0.025 | 0.024 | 0.025 | 0.0288 | 0.0236 | 0.025522 | 0.024 | 0.024867 | 0.0245 | 0.0188 | 0.02 | 0.0236 | 0.0239 | 0.03 | 0.023 | 0.023 |  |  | 0.024514
| 1 | 5 | 0.0237 | 0.027 | 0.02416 | 0.024 | 0.02 | 0.026 | 0.025 | 0.027 | 0.024 | 0.025 | 0.024 | 0.029 | 0.0234 | 0.026083 | 0.023 | 0.024433 | 0.0246 | 0.0255 | 0.023 | 0.0243 | 0.0247 | 0.03 | 0.024 | 0.024 |  |  | 153
| 1 | 6 | 0.0238 | 0.0273 | 0.02351 | 0.024 | 0.02 | 0.025 | 0.025 | 0.027 | 0.025 | 0.025 | 0.024 | 0.029 | 0.0236 | 0.025875 | 0.024 | 0.0248 | 0.0246 | 0.0221 | 0.018 | 0.0236 | 0.0241 | 0.03 | 0.023 | 0.023
| Mean |  | 0.023783 | 0.026967 | 0.023591 | 0.024 | 0.02 | 0.025333 | 0.025333 | 0.026667 | 0.0245 | 0.024333 | 0.0245 | 0.028583 | 0.024067 | 0.025619 | 0.023333 | 0.024667 | 0.024567 | 0.023533 | 0.020833 | 0.024 | 0.0242 | 0.03 | 0.023167 | 0.023167
| Median |  | 0.0238 | 0.02695 | 0.023416 | 0.024 | 0.02 | 0.025 | 0.0255 | 0.027 | 0.0245 | 0.024 | 0.0245 | 0.0289 | 0.02385 | 0.02553 | 0.023 | 0.02465 | 0.0246 | 0.0233 | 0.021 | 0.024 | 0.0241 | 0.03 | 0.023 | 0.023
| Std Dev. |  | 0.00013292 | 0.0002582 | 0.00060543 | 0.00063246 | 0 | 0.0005164 | 0.0008165 | 0.0005164 | 0.00054772 | 0.0005164 | 0.00054772 | 0.00058109 | 0.00065625 | 0.00029839 | 0.0005164 | 0.00023878 | 0.00010328 | 0.0032586 | 0.0017224 | 0.00040497 | 0.00046476 | 0 | 0.00040825 | 0.00040825
| Rel.Std.Dev. |  | 0.5589% | 0.9575% | 2.566% | 2.635% | 0% | 2.038% | 3.223% | 1.936% | 2.236% | 2.122% | 2.236% | 2.033% | 2.727% | 1.165% | 2.213% | 0.968% | 0.4204% | 13.85% | 8.268% | 1.687% | 1.92% | 0% | 1.762% | 1.762%
| PDM3 |  | -2.979% | 10.01% | -3.762% | -2.095% | -18.41% | 3.344% | 3.344% | 8.783% | -0.05577% | -0.7357% | -0.05577% | 16.6% | -1.823% | 4.508% | -4.815% | 0.6241% | 0.2162% | -3.999% | -15.01% | -2.095% | -1.28% | 22.38% | -5.495% | -5.495%
| Z-Score (Absolute) |  | 0.39 | 2.18 | 0.55 | 0.22 | 3.45 | 0.86 | 0.86 | 1.94 | 0.19 | 0.05 | 0.19 | 3.49 | 0.16 | 1.09 | 0.76 | 0.32 | 0.24 | 0.59 | 2.78 | 0.22 | 0.05 | 4.64 | 0.89 | 0.89
| Analytical results for Pb in OREAS 906b (Certified Value 16.9 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 2
| 1 | 1 | 16.6 | 17.9 | 16.0255 | 14.01 | 24.1 | 15.6 | 16.9 | 17.5 | 15.8 | 16 | 17 | 43 | 16.5 | 17.2766 | 16 | 16.8 | 17.19 | 17 | 17.3 | 16.1 | 23.4 | 19 | 17.4 | 18 | 15.94 |  | 1
| 1 | 2 | 17 | 17.9 | 16.3287 | 15.55 | 23.7 | 16.6 | 16.6 | 17.4 | 16.1 | 16.4 | 16.8 | 20 | 16.2 | 18.2553 | 17 | 16.6 | 17.73 | 16 | 18.2 | 16.7 | 22.3 | 18.1 | 17.2 | 19 | 16.54 |  | 29
| 1 | 3 | 16.4 | 17.9 | 16.9218 | 15.22 | 23.7 | 17 | 16.6 | 17.2 | 17.7 | 16.8 | 17.7 | 19 | 16.8 | 17.2085 | 17 | 16.3 | 17.69 | 17 | 17.5 | 16.9 | 22.6 | 18.6 | 17.5 | 18 | 16.4 |  | 16
| 1 | 4 | 17.1 | 18.4 | 16.5255 | 14.54 | 24.2 | 16.4 | 15.9 | 17.1 | 16 | 15.8 | 17.4 | 21 | 15.9 | 17.6413 | 16 | 16.7 | 16.55 | 16 | 17.2 | 15.9 | 22.9 | 18.6 | 17.3 | 20 | 16.02 |  | 16.8719
| 1 | 5 | 16.3 | 18.4 | 17.319 | 15.32 | 23.5 | 16.2 | 16.8 | 17.1 | 15.8 | 16.4 | 16.8 | 19 | 16.8 | 18.2005 | 17 | 16 | 17.44 | 17 | 17.4 | 16.6 | 23 | 19.2 | 17.4 | 19 | 16.7 |  | 154
| 1 | 6 | 16.6 | 18.3 | 16.8206 | 14.2 | 24.1 | 16.2 | 16.9 | 17.8 | 16.4 | 15.9 | 17.4 | 24 | 16.4 | 18.9632 | 17 | 16 | 17.8 | 18 | 17.2 | 16.5 | 22.4 | 18.6 | 17.3 | 18 | 16.31
| Mean |  | 16.6667 | 18.1333 | 16.6569 | 14.8067 | 23.8833 | 16.3333 | 16.6167 | 17.35 | 16.3 | 16.2167 | 17.1833 | 24.3333 | 16.4333 | 17.9242 | 16.6667 | 16.4 | 17.4 | 16.8333 | 17.4667 | 16.45 | 22.7667 | 18.6833 | 17.35 | 18.6667 | 16.3183
| Median |  | 16.6 | 18.1 | 16.673 | 14.88 | 23.9 | 16.3 | 16.7 | 17.3 | 16.05 | 16.2 | 17.2 | 20.5 | 16.45 | 17.9209 | 17 | 16.45 | 17.565 | 17 | 17.35 | 16.55 | 22.75 | 18.6 | 17.35 | 18.5 | 16.355
| Std Dev. |  | 0.32042 | 0.2582 | 0.46019 | 0.64199 | 0.28577 | 0.46762 | 0.37639 | 0.27386 | 0.72111 | 0.38166 | 0.37103 | 9.331 | 0.35024 | 0.67485 | 0.5164 | 0.35214 | 0.47354 | 0.75277 | 0.37771 | 0.37815 | 0.41312 | 0.38166 | 0.10488 | 0.8165 | 0.29451
| Rel.Std.Dev. |  | 1.922% | 1.424% | 2.763% | 4.336% | 1.197% | 2.863% | 2.265% | 1.578% | 4.424% | 2.354% | 2.159% | 38.35% | 2.131% | 3.765% | 3.098% | 2.147% | 2.721% | 4.472% | 2.162% | 2.299% | 1.815% | 2.043% | 0.6045% | 4.374% | 1.805%
| PDM3 |  | -1.216% | 7.477% | -1.274% | -12.24% | 41.56% | -3.192% | -1.513% | 2.834% | -3.389% | -3.883% | 1.846% | 44.22% | -2.599% | 6.237% | -1.216% | -2.797% | 3.13% | -0.2284% | 3.525% | -2.5% | 34.94% | 10.74% | 2.834% | 10.64% | -3.281%
| Z-Score (Absolute) |  | 0 | 1.91 | 0.01 | 2.43 | 9.42 | 0.44 | 0.07 | 0.89 | 0.48 | 0.59 | 0.67 | NA | 0.3 | 1.64 | NA | 0.35 | 0.96 | NA | 1.04 | 0.28 | 7.96 | 2.63 | 0.89 | NA | 0.45
Results from laboratories 2.13, 2.16, 2.19 and 2.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 906b (Indicative Value < 10 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | [25 empty] | 1
| No | No | 2.03 | [25 empty] | ppb
|  |  | AR*OES/MS | [25 empty] | 1
|  |  | 01g | [25 empty] | 1
| 1 | 1 | < 10 | [25 empty] | 1
| 1 | 2 | < 10 | [25 empty] | 1
| 1 | 3 | < 10 | [25 empty] | 16
| 1 | 4 | < 10 | [25 empty] | < 10
| 1 | 5 | < 10 | [25 empty] | 25
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 906b (Indicative Value 8.68 ppm) | [27 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 8.522 | 7.9034 | 9.3198 | 9.3 | [22 empty] | 1
| 1 | 2 | 8.396 | 7.7945 | 9.3467 | 9.2 | [22 empty] | 20
| 1 | 3 | 8.338 | 7.9558 | 9.2339 | 9.2 | [22 empty] | 16
| 1 | 4 | 8.614 | 7.9119 | 9.3778 | 9.1 | [22 empty] | 8.6793
| 1 | 5 | 8.495 | 7.8892 | 9.1069 | 8.9 | [22 empty] | 26
| 1 | 6 | 8.455 | 7.8169 | 9.0262 | 9.1
| Mean |  | 8.47 | 7.8786 | 9.2352 | 9.1333
| Median |  | 8.475 | 7.8963 | 9.2769 | 9.15
| Std Dev. |  | 0.097149 | 0.061111 | 0.14145 | 0.13663
| Rel.Std.Dev. |  | 1.147% | 0.7757% | 1.532% | 1.496%
| PDM3 |  | -2.411% | -9.225% | 6.405% | 5.231%
| Z-Score (Absolute) |  | 0.58 | 1.63 | 0.76 | 0.58
| Analytical results for Pt in OREAS 906b (Indicative Value 3.33 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 2.03 | 2.20 | [24 empty] | ppb
|  |  | AR*OES/MS | AR*OES/MS | [24 empty] | 2
|  |  | 01g | 0.1g | [24 empty] | 2
| 1 | 1 | < 5 | 4 | [24 empty] | 1
| 1 | 2 | < 5 | 3 | [24 empty] | 1
| 1 | 3 | < 5 | 3 | [24 empty] | 16
| 1 | 4 | < 5 | 2 | [24 empty] | 3.3333
| 1 | 5 | < 5 | 4 | [24 empty] | 22
| 1 | 6 | < 5 | 4
| Mean |  |  | 3.3333
| Median |  |  | 3.5
| Std Dev. |  |  | 0.8165
| Rel.Std.Dev. |  |  | 24.49%
| PDM3 |  |  | 1.11e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 906b (Certified Value 12.6 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 1
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [8 empty] | 1
| 1 | 1 | 13.05 | 17.4 | 12.6561 | 61.9 | 10.2 | 12.2 | 13.4 | 10.4 | 11.5 | 12.27 | 11.7 | 13.4762 | 11.6 | 30.8 | 15.32 | 11.5 | 12.7 | 14.6 | [8 empty] | 1
| 1 | 2 | 12.87 | 17.6 | 13.1739 | 62.4 | 10.3 | 12.1 | 13 | 11.1 | 11.6 | 12.02 | 11.9 | 13.2085 | 11.3 | 30 | 14.01 | 12.3 | 12.9 | 15.5 | [8 empty] | 45
| 1 | 3 | 12.73 | 17.4 | 13.8296 | 60.3 | 10.9 | 12.1 | 13.2 | 10.6 | 11.5 | 12.81 | 11.8 | 13.4131 | 11.3 | 29.1 | 15.2 | 11.7 | 13 | 15.5 | [8 empty] | 16
| 1 | 4 | 13.29 | 18.3 | 13.057 | 57.9 | 10.5 | 11.4 | 13.2 | 10.5 | 11.1 | 13.68 | 11.8 | 13.4689 | 11.9 | 28.7 | 15.39 | 11.7 | 12.7 | 14.9 | [8 empty] | 12.5692
| 1 | 5 | 12.56 | 18 | 13.8088 | 59.4 | 10.5 | 12.2 | 13.6 | 10.3 | 11.6 | 13.09 | 12 | 14.0029 | 12 | 26.4 | 14.88 | 12.1 | 13 | 14.3 | [8 empty] | 155
| 1 | 6 | 13.03 | 17.7 | 13.297 | 60 | 10.2 | 11.4 | 13.6 | 10.8 | 11.4 | 13.47 | 11.5 | 12.8831 | 11.8 | 31.9 | 14.66 | 12.2 | 13.3 | 15.3
| Mean |  | 12.9217 | 17.7333 | 13.3037 | 60.3167 | 10.4333 | 11.9 | 13.3333 | 10.6167 | 11.45 | 12.89 | 11.7833 | 13.4088 | 11.65 | 29.4833 | 14.91 | 11.9167 | 12.9333 | 15.0167
| Median |  | 12.95 | 17.65 | 13.2355 | 60.15 | 10.4 | 12.1 | 13.3 | 10.55 | 11.5 | 12.95 | 11.8 | 13.441 | 11.7 | 29.55 | 15.04 | 11.9 | 12.95 | 15.1
| Std Dev. |  | 0.25849 | 0.3559 | 0.45362 | 1.651 | 0.26583 | 0.38987 | 0.24221 | 0.29269 | 0.18708 | 0.65547 | 0.17224 | 0.36813 | 0.30166 | 1.9031 | 0.52077 | 0.32506 | 0.22509 | 0.49967
| Rel.Std.Dev. |  | 2% | 2.007% | 3.41% | 2.737% | 2.548% | 3.276% | 1.817% | 2.757% | 1.634% | 5.085% | 1.462% | 2.745% | 2.589% | 6.455% | 3.493% | 2.728% | 1.74% | 3.327%
| PDM3 |  | 2.804% | 41.09% | 5.844% | 379.9% | -16.99% | -5.324% | 6.08% | -15.53% | -8.904% | 2.553% | -6.252% | 6.68% | -7.313% | 134.6% | 18.62% | -5.191% | 2.897% | 19.47%
| Z-Score (Absolute) |  | 0 | 2.68 | 0.21 | 26.39 | 1.39 | 0.57 | 0.23 | 1.29 | 0.82 | 0.02 | 0.64 | 0.27 | 0.71 | 9.22 | 1.1 | 0.56 | 0 | 1.16
| Analytical results for Re in OREAS 906b (Indicative Value 0.001 ppm) | [27 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.20 | 2.22 | 2.24 | 2.25 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [12 empty] | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.1g | 0.5g | 0.25g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.001 | < 0.05 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.001 | < 0.02 | < 0.001 | [12 empty] | 1
| 1 | 2 |  | < 0.05 | < 0.05 | 0.001 | 0.002 | 0.001 | 0.001 | 0.002 | 0.003 | < 0.05 | 0.001 | 0.001 | < 0.02 | < 0.001 | [12 empty] | 12
| 1 | 3 |  | < 0.05 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.003 | < 0.05 | 0.002 | < 0.001 | < 0.02 | < 0.001 | [12 empty] | 16
| 1 | 4 | 0.002 | < 0.05 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.002 | < 0.05 | 0.001 | 0.001 | < 0.02 | 0.001 | [12 empty] | 0.0012593
| 1 | 5 | 0.002 | < 0.05 | < 0.05 | 0.001 | 0.002 | 0.001 | 0.001 | 0.002 | 0.003 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | [12 empty] | 23
| 1 | 6 | 0.001 | < 0.05 | < 0.05 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | 0.003 | < 0.05 | 0.001 | 0.002 | < 0.02 | < 0.001
| Mean |  | 0.0015 |  |  | 0.0011667 | 0.0013333 | 0.001 | 0.001 | 0.0016667 | 0.0026667 |  | 0.0013333 | 0.0013333 |  | 0.001
| Median |  | 0.0015 |  |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.003 |  | 0.001 | 0.001 |  | 0.001
| Std Dev. |  | 0.00057735 |  |  | 0.00040825 | 0.0005164 | 0 | 0 | 0.0005164 | 0.0005164 |  | 0.0005164 | 0.00057735
| Rel.Std.Dev. |  | 38.49% |  |  | 34.99% | 38.73% | 0% | 0% | 30.98% | 19.36% |  | 38.73% | 43.3%
| PDM3 |  | 19.12% |  |  | -7.353% | 5.882% | -20.59% | -20.59% | 32.35% | 111.8% |  | 5.882% | 5.882% |  | -20.59%
| Z-Score (Absolute) |  | 0.12 | 34.7 | 34.7 | 0.37 | 0.12 | 0.61 | 0.61 | 0.37 | 1.84 | 34.7 | 0.12 | 0.74 | 12.63 | 1.23
| Analytical results for S in OREAS 906b (Certified Value 0.04 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 3
| 1 | 1 | 0.0408 | 0.04 | 0.041456 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.07 | 0.05 | 0.04 | 0.046 | 0.04 | 0.041593 | 0.04 | 0.04 | 0.036049 | 0.04 | 0.0399 | 0.04 | 0.05 | 0.0431 | 0.04 | 0.039 | 0.0431 | 0.04 | 1
| 1 | 2 | 0.0424 | 0.04 | 0.038693 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.08 | 0.06 | 0.04 | 0.04 | 0.04 | 0.040875 | 0.04 | 0.04 | 0.036086 | 0.04 | 0.0463 | 0.04 | 0.05 | 0.042 | 0.05 | 0.04 | 0.0428 | 0.04 | 3
| 1 | 3 | 0.0422 | 0.04 | 0.03915 | 0.04 | 0.04 | 0.06 | 0.05 | 0.04 | 0.07 | 0.05 | 0.04 | 0.04 | 0.04 | 0.041445 | 0.04 | 0.04 | 0.036935 | 0.04 | 0.0436 | 0.04 | 0.05 | 0.0427 | 0.04 | 0.04 | 0.0423 | 0.04 | 16
| 1 | 4 | 0.0423 | 0.04 | 0.041454 | 0.04 | 0.04 | 0.06 | 0.05 | 0.04 | 0.07 | 0.05 | 0.04 | 0.04 | 0.04 | 0.041178 | 0.04 | 0.04 | 0.036366 | 0.04 | 0.0448 | 0.04 | 0.05 | 0.0442 | 0.04 | 0.039 | 0.0428 | 0.04 | 0.040411
| 1 | 5 | 0.0414 | 0.04 | 0.042129 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.07 | 0.06 | 0.04 | 0.04 | 0.04 | 0.040989 | 0.04 | 0.04 | 0.036281 | 0.04 | 0.0474 | 0.04 | 0.05 | 0.0439 | 0.05 | 0.04 | 0.0438 | 0.03 | 156
| 1 | 6 | 0.0416 | 0.04 | 0.039194 | 0.04 | 0.04 | 0.04 | 0.05 | 0.04 | 0.07 | 0.06 | 0.04 | 0.04 | 0.04 | 0.04129 | 0.04 | 0.04 | 0.036377 | 0.04 | 0.0371 | 0.03 | 0.05 | 0.0434 | 0.04 | 0.04 | 0.0423 | 0.04
| Mean |  | 0.041783 | 0.04 | 0.040346 | 0.04 | 0.04 | 0.051667 | 0.05 | 0.04 | 0.071667 | 0.055 | 0.04 | 0.041 | 0.04 | 0.041228 | 0.04 | 0.04 | 0.036349 | 0.04 | 0.043183 | 0.038333 | 0.05 | 0.043217 | 0.043333 | 0.039667 | 0.04285 | 0.038333
| Median |  | 0.0419 | 0.04 | 0.040324 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.07 | 0.055 | 0.04 | 0.04 | 0.04 | 0.041234 | 0.04 | 0.04 | 0.036323 | 0.04 | 0.0442 | 0.04 | 0.05 | 0.04325 | 0.04 | 0.04 | 0.0428 | 0.04
| Std Dev. |  | 0.00062743 | 0 | 0.0014918 | 0 | 0 | 0.0075277 | 7.6012e-18 | 0 | 0.0040825 | 0.0054772 | 0 | 0.0024495 | 0 | 0.00027165 | 0 | 0 | 0.00031897 | 0 | 0.0039514 | 0.0040825 | 7.6012e-18 | 0.00080353 | 0.005164 | 0.0005164 | 0.00056125 | 0.0040825
| Rel.Std.Dev. |  | 1.502% | 0% | 3.697% | 0% | 0% | 14.57% | 1.52e-14% | 0% | 5.696% | 9.959% | 0% | 5.974% | 0% | 0.6589% | 0% | 0% | 0.8775% | 0% | 9.15% | 10.65% | 1.52e-14% | 1.859% | 11.92% | 1.302% | 1.31% | 10.65%
| PDM3 |  | 3.397% | -1.016% | -0.1598% | -1.016% | -1.016% | 27.85% | 23.73% | -1.016% | 77.35% | 36.1% | -1.016% | 1.458% | -1.016% | 2.024% | -1.016% | -1.016% | -10.05% | -1.016% | 6.861% | -5.14% | 23.73% | 6.944% | 7.232% | -1.841% | 6.036% | -5.14%
| Z-Score (Absolute) |  | 0.81 | 0.09 | 0.09 | 0.09 | 0.09 | 5.81 | 4.97 | 0.09 | 15.93 | 7.5 | 0.09 | 0.42 | 0.09 | 0.53 | 0.09 | 0.09 | 1.93 | 0.09 | 1.52 | 0.93 | 4.97 | 1.54 | 1.6 | 0.26 | 1.35 | 0.93
| Analytical results for Sb in OREAS 906b (Certified Value 1.26 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 2
| 1 | 1 | 1.5 | 1.33 | 1.0542 | 1.18 | < 5 | 1.14 | 1.2 | 1.35 | 1.02 | 1.16 | 0.83 | < 2 | 1.24 | 1.172 | < 5 | 1.36 | < 5 | 1.28 | 1.73 | 1.23 | 1.395 | 1.26 | 1.45 | 1.31 | 1.07 |  | 1
| 1 | 2 | 1.51 | 1.32 | 1.1139 | 1.36 | < 5 | 1.14 | 1.22 | 1.37 | 1.12 | 1.18 | 0.86 | < 2 | 1.26 | 1.0998 | < 5 | 1.34 | < 5 | 1.3 | 1.63 | 1.28 | 1.55 | 1.28 | 1.43 | 1.34 | 1.07 |  | 47
| 1 | 3 | 1.47 | 1.34 | 1.1659 | 1.27 | < 5 | 1.18 | 1.26 | 1.34 | 1.14 | 1.14 | 0.86 | < 2 | 1.23 | 1.1795 | < 5 | 1.31 | < 5 | 1.29 | 1.7 | 1.22 | 1.485 | 1.26 | 1.48 | 1.26 | 1.06 |  | 16
| 1 | 4 | 1.46 | 1.33 | 1.0285 | 1.17 | < 5 | 1.12 | 1.36 | 1.31 | 1.03 | 1.16 | 0.86 | < 2 | 1.25 | 1.1184 | < 5 | 1.31 | < 5 | 1.32 | 1.64 | 1.19 | 1.44 | 1.28 | 1.46 | 1.35 | 1.09 |  | 1.263
| 1 | 5 | 1.44 | 1.32 | 1.1452 | 1.22 | < 5 | 1.16 | 1.24 | 1.38 | 1.04 | 1.18 | 0.84 | < 2 | 1.22 | 1.1044 | < 5 | 1.37 | < 5 | 1.38 | 1.73 | 1.28 | 1.54 | 1.24 | 1.45 | 1.48 | 1.13 |  | 157
| 1 | 6 | 1.48 | 1.33 | 1.1514 | 1.32 | < 5 | 1.27 | 1.22 | 1.34 | 1.04 | 1.13 | 0.86 | < 2 | 1.2 | 1.0916 | < 5 | 1.33 | < 5 | 1.29 | 1.6 | 1.25 | 1.54 | 1.28 | 1.46 | 1.42 | 1.07
| Mean |  | 1.4767 | 1.3283 | 1.1099 | 1.2533 |  | 1.1683 | 1.25 | 1.3483 | 1.065 | 1.1583 | 0.85167 |  | 1.2333 | 1.1276 |  | 1.3367 |  | 1.31 | 1.6717 | 1.2417 | 1.4917 | 1.2667 | 1.455 | 1.36 | 1.0817
| Median |  | 1.475 | 1.33 | 1.1296 | 1.245 |  | 1.15 | 1.23 | 1.345 | 1.04 | 1.16 | 0.86 |  | 1.235 | 1.1114 |  | 1.335 |  | 1.295 | 1.67 | 1.24 | 1.5125 | 1.27 | 1.455 | 1.345 | 1.07
| Std Dev. |  | 0.02582 | 0.0075277 | 0.056303 | 0.076855 |  | 0.053821 | 0.057619 | 0.024833 | 0.051284 | 0.020412 | 0.013292 |  | 0.021602 | 0.038358 |  | 0.025033 |  | 0.036878 | 0.055648 | 0.035449 | 0.063456 | 0.01633 | 0.016432 | 0.07874 | 0.025626
| Rel.Std.Dev. |  | 1.749% | 0.5667% | 5.073% | 6.132% |  | 4.607% | 4.61% | 1.842% | 4.815% | 1.762% | 1.561% |  | 1.752% | 3.402% |  | 1.873% |  | 2.815% | 3.329% | 2.855% | 4.254% | 1.289% | 1.129% | 5.79% | 2.369%
| PDM3 |  | 16.92% | 5.175% | -12.12% | -0.7632% |  | -7.493% | -1.027% | 6.759% | -15.68% | -8.285% | -32.57% |  | -2.347% | -10.72% |  | 5.835% |  | 3.724% | 32.36% | -1.687% | 18.11% | 0.2925% | 15.2% | 7.682% | -14.36%
| Z-Score (Absolute) |  | 1.02 | 0.3 | 0.76 | 0.06 | 5.98 | 0.48 | 0.08 | 0.4 | 0.98 | 0.53 | 2.01 | 1.29 | 0.16 | 0.67 | 5.98 | 0.34 | 5.98 | 0.21 | 1.96 | 0.12 | 1.09 | 0 | 0.91 | 0.45 | 0.9
| Analytical results for Sc in OREAS 906b (Certified Value 1.39 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 2
| 1 | 1 | 1.2 | 1 | 1.2288 | 1.4 | < 1 | 1.2 | 1.4 | 1.2 | 1.2 | 1.4 | 1.6 | 1.3 | 1.4 | 1.3066 | 0.8 | 1.5 | < 5 | 2 | 1.5 | 1.5 | 1.7 | 1.6 | 1.4 | 1.2 |  |  | 1
| 1 | 2 | 1.2 | 1 | 1.3628 | 1.5 | < 1 | 1.2 | 1.4 | 1.2 | 1.3 | 1.4 | 1.6 | 1 | 1.4 | 1.3272 | 0.8 | 1.4 | < 5 | 1 | 1.4 | 1.5 | 1.8 | 1.6 | 1.4 | 1.2 |  |  | 14
| 1 | 3 | 1.2 | 1 | 1.4351 | 1.5 | < 1 | 1.3 | 1.4 | 1.2 | 1.3 | 1.4 | 1.6 | 1 | 1.4 | 1.4176 | 0.8 | 1.5 | < 5 | 1 | 1.7 | 1.5 | 1.7 | 1.6 | 1.6 | 1.2 |  |  | 16
| 1 | 4 | 1.2 | 1 | 1.2422 | 1.3 | < 1 | 1.2 | 1.3 | 1.2 | 1.3 | 1.3 | 1.6 | 1.2 | 1.3 | 1.3925 | 0.8 | 1.5 | < 5 | 2 | 1.7 | 1.6 | 1.8 | 1.7 | 1.5 | 1.2 |  |  | 1.3916
| 1 | 5 | 1.2 | 1 | 1.382 | 1.3 | < 1 | 1.3 | 1.4 | 1.2 | 1.3 | 1.4 | 1.6 | 1 | 1.3 | 1.3896 | 0.8 | 1.5 | < 5 | 2 | 1.6 | 1.5 | 1.8 | 1.6 | 1.4 | 1.2 |  |  | 158
| 1 | 6 | 1.2 | 1 | 1.3879 | 1.5 | < 1 | 1.2 | 1.4 | 1.2 | 1.3 | 1.4 | 1.6 | 1 | 1.3 | 1.3302 | 0.8 | 1.4 | < 5 | 1 | 1.6 | 1.5 | 1.8 | 1.7 | 1.6 | 1.2
| Mean |  | 1.2 | 1 | 1.3398 | 1.4167 |  | 1.2333 | 1.3833 | 1.2 | 1.2833 | 1.3833 | 1.6 | 1.0833 | 1.35 | 1.3606 | 0.8 | 1.4667 |  | 1.5 | 1.5833 | 1.5167 | 1.7667 | 1.6333 | 1.4833 | 1.2
| Median |  | 1.2 | 1 | 1.3724 | 1.45 |  | 1.2 | 1.4 | 1.2 | 1.3 | 1.4 | 1.6 | 1 | 1.35 | 1.3599 | 0.8 | 1.5 |  | 1.5 | 1.6 | 1.5 | 1.8 | 1.6 | 1.45 | 1.2
| Std Dev. |  | 0 | 0 | 0.084331 | 0.098319 |  | 0.05164 | 0.040825 | 0 | 0.040825 | 0.040825 | 2.4324e-16 | 0.13292 | 0.054772 | 0.044842 | 1.2162e-16 | 0.05164 |  | 0.54772 | 0.1169 | 0.040825 | 0.05164 | 0.05164 | 0.098319 | 0
| Rel.Std.Dev. |  | 0% | 0% | 6.294% | 6.94% |  | 4.187% | 2.951% | 0% | 3.181% | 2.951% | 1.52e-14% | 12.27% | 4.057% | 3.296% | 1.52e-14% | 3.521% |  | 36.51% | 7.383% | 2.692% | 2.923% | 3.162% | 6.628% | 0%
| PDM3 |  | -13.77% | -28.14% | -3.722% | 1.801% |  | -11.37% | -0.5941% | -13.77% | -7.78% | -0.5941% | 14.98% | -22.15% | -2.989% | -2.227% | -42.51% | 5.394% |  | 7.79% | 13.78% | 8.987% | 26.95% | 17.37% | 6.592% | -13.77%
| Z-Score (Absolute) |  | 0.73 | NA | 0.14 | 0.19 | 3.71 | 0.59 | 0.05 | 0.73 | 0.38 | 0.05 | 0.97 | 1.23 | 0.09 | 0.05 | 2.44 | 0.4 | 4.8 | NA | 0.9 | 0.62 | 1.68 | 1.11 | 0.47 | 0.73
Results from laboratories 2.02 and 2.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 906b (Certified Value 4.52 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  |  |  | 3
| 1 | 1 | 4 | 5 | 4.1462 | 3.8 | < 5 | 4.3 | 4.3 | 4.6 | 4.6 | 4.2 | 5 | 4.906 | 4.7 | < 10 | 4.76 | 8 | 3 | 4.3 | 5 | 4.3 | 4.96 | 4.6 |  |  |  |  | 1
| 1 | 2 | 4 | 5 | 4.011 | 4.1 | < 5 | 4.6 | 4.8 | 4.5 | 4.5 | 4.3 | 5 | 4.3213 | 4.6 | < 10 | 4.72 | 6 | 3 | 4.1 | 5 | 4.1 | 4.85 | 4.5 |  |  |  |  | 49
| 1 | 3 | 4 | 5 | 4.7982 | 3.2 | < 5 | 4.8 | 4.4 | 4.5 | 4.6 | 4.3 | 5 | 4.6235 | 4.2 | < 10 | 5.18 | 6 | 3 | 4.2 | 5 | 4.5 | 4.89 | 4.4 |  |  |  |  | 16
| 1 | 4 | 4 | 5 | 4.5501 | 3.6 | < 5 | 4.2 | 4.5 | 4.4 | 4.6 | 4.1 | 5 | 5.041 | 4.7 | < 10 | 4.72 | 8 | 3 | 4.4 | 4 | 4.4 | 5 | 4.6 |  |  |  |  | 4.5199
| 1 | 5 | 4 | 5 | 4.7196 | 4.2 | < 5 | 4.7 | 4.2 | 4.8 | 4.1 | 4.2 | 5 | 4.6131 | 4.2 | < 10 | 4.95 | 6 | 3 | 4.3 | 5 | 4.4 | 4.97 | 4.7 |  |  |  |  | 159
| 1 | 6 | 4 | 5 | 4.7689 | 3.5 | < 5 | 3.9 | 4.4 | 4.2 | 4.9 | 4.4 | 4 | 4.5331 | 4.3 | < 10 | 4.92 | 7 | 3 | 4.4 | 5 | 4.4 | 4.82 | 4.3
| Mean |  | 4 | 5 | 4.499 | 3.7333 |  | 4.4167 | 4.4333 | 4.5 | 4.55 | 4.25 | 4.8333 | 4.673 | 4.45 |  | 4.875 | 6.8333 | 3 | 4.2833 | 4.8333 | 4.35 | 4.915 | 4.5167
| Median |  | 4 | 5 | 4.6349 | 3.7 |  | 4.45 | 4.4 | 4.5 | 4.6 | 4.25 | 5 | 4.6183 | 4.45 |  | 4.84 | 6.5 | 3 | 4.3 | 5 | 4.4 | 4.925 | 4.55
| Std Dev. |  | 0 | 0 | 0.33951 | 0.37771 |  | 0.34303 | 0.20656 | 0.2 | 0.25884 | 0.10488 | 0.40825 | 0.26039 | 0.2429 |  | 0.17997 | 0.98319 | 0 | 0.1169 | 0.40825 | 0.13784 | 0.072319 | 0.1472
| Rel.Std.Dev. |  | 0% | 0% | 7.546% | 10.12% |  | 7.767% | 4.659% | 4.444% | 5.689% | 2.468% | 8.447% | 5.572% | 5.458% |  | 3.692% | 14.39% | 0% | 2.729% | 8.447% | 3.169% | 1.471% | 3.259%
| PDM3 |  | -11.5% | 10.62% | -0.4628% | -17.4% |  | -2.285% | -1.916% | -0.4408% | 0.6654% | -5.972% | 6.934% | 3.386% | -1.547% |  | 7.856% | 51.18% | -33.63% | -5.234% | 6.934% | -3.759% | 8.741% | -0.07211%
| Z-Score (Absolute) |  | 1.15 | 1.16 | 0 | 1.77 | 4.62 | 0.19 | 0.15 | 0 | 0.12 | 0.58 | 0.77 | 0.4 | 0.11 | 1.16 | 0.87 | 5.4 | 3.47 | 0.5 | 0.77 | 0.35 | 0.96 | 0.04
| Analytical results for Sm in OREAS 906b (Indicative Value 5.34 ppm) | [27 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 2.01 | 2.15 | 2.18 | 2.25 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [22 empty] | 2
| 1 | 1 | 5.044 | 4.7493 | 5.8565 | 5.7 | [22 empty] | 1
| 1 | 2 | 5.077 | 4.8594 | 5.7763 | 5.8 | [22 empty] | 18
| 1 | 3 | 5.027 | 4.8237 | 5.935 | 6 | [22 empty] | 16
| 1 | 4 | 5.188 | 4.6175 | 5.8334 | 5.6 | [22 empty] | 5.3405
| 1 | 5 | 4.879 | 4.735 | 5.6154 | 5.8 | [22 empty] | 24
| 1 | 6 | 5.151 | 4.7153 | 5.4898 | 5.9
| Mean |  | 5.061 | 4.75 | 5.7511 | 5.8
| Median |  | 5.0605 | 4.7421 | 5.8048 | 5.8
| Std Dev. |  | 0.10868 | 0.085312 | 0.16674 | 0.14142
| Rel.Std.Dev. |  | 2.147% | 1.796% | 2.899% | 2.438%
| PDM3 |  | -5.234% | -11.06% | 7.687% | 8.604%
| Z-Score (Absolute) |  | 0.63 | 1.2 | 0.63 | 0.72
| Analytical results for Sn in OREAS 906b (Certified Value 1.32 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 1.39 | 1.4 | 1.31 | < 5 | 1.2 | 1.3 | 1.3 | 1.3 | 1.2 | 1.3 | 1.3 | 1.2967 | < 10 | 1.4 | < 20 | 1.9 | 2.1 | 0.7 | 1.3 | 1.58 | [6 empty] | 1
| 1 | 2 | 1.39 | 1.4 | 1.3033 | < 5 | 1.3 | 1.3 | 1.3 | 1.3 | 1.3 | 1.3 | 1.3 | 1.2466 | < 10 | 1.3 | < 20 | 1.9 | 1.9 | 0.7 | 1.3 | 1.56 | [6 empty] | 10
| 1 | 3 | 1.36 | 1.4 | 1.4085 | < 5 | 1.3 | 1.2 | 1.3 | 1.3 | 1.3 | 1.3 | 1.5 | 1.2631 | < 10 | 1.2 | < 20 | 1.8 | 2.2 | 0.7 | 1.3 | 1.55 | [6 empty] | 16
| 1 | 4 | 1.4 | 1.7 | 1.3126 | < 5 | 1.2 | 1.3 | 1.3 | 1.3 | 1.3 | 1.4 | 1.3 | 1.2748 | < 10 | 1.3 | < 20 | 1.9 | 2 | 0.8 | 1.3 | 1.56 | [6 empty] | 1.3193
| 1 | 5 | 1.38 | 1.4 | 1.3833 | < 5 | 1.3 | 1.3 | 1.4 | 1.3 | 1.2 | 1.4 | 1.5 | 1.262 | < 10 | 1.3 | < 20 | 1.8 | 2.2 | 0.7 | 1.3 | 1.74 | [6 empty] | 160
| 1 | 6 | 1.37 | 1.6 | 1.4 | < 5 | 1.3 | 1.3 | 1.4 | 1.3 | 1.2 | 1.4 | 1.3 | 1.256 | < 10 | 1.3 | < 20 | 1.9 | 2.1 | 0.7 | 1.3 | 1.55
| Mean |  | 1.3817 | 1.4833 | 1.3529 |  | 1.2667 | 1.2833 | 1.3333 | 1.3 | 1.25 | 1.35 | 1.3667 | 1.2665 |  | 1.3 |  | 1.8667 | 2.0833 | 0.71667 | 1.3 | 1.59
| Median |  | 1.385 | 1.4 | 1.348 |  | 1.3 | 1.3 | 1.3 | 1.3 | 1.25 | 1.35 | 1.3 | 1.2625 |  | 1.3 |  | 1.9 | 2.1 | 0.7 | 1.3 | 1.56
| Std Dev. |  | 0.01472 | 0.13292 | 0.049322 |  | 0.05164 | 0.040825 | 0.05164 | 0 | 0.054772 | 0.054772 | 0.10328 | 0.017441 |  | 0.063246 |  | 0.05164 | 0.1169 | 0.040825 | 0 | 0.074297
| Rel.Std.Dev. |  | 1.065% | 8.961% | 3.646% |  | 4.077% | 3.181% | 3.873% | 0% | 4.382% | 4.057% | 7.557% | 1.377% |  | 4.865% |  | 2.766% | 5.611% | 5.696% | 0% | 4.673%
| PDM3 |  | 4.726% | 12.43% | 2.549% |  | -3.991% | -2.728% | 1.062% | -1.464% | -5.254% | 2.325% | 3.589% | -4% |  | -1.464% |  | 41.49% | 57.91% | -45.68% | -1.464% | 20.52%
| Z-Score (Absolute) |  | 0.24 | 1.05 | 0.01 | 9.13 | 0.67 | 0.54 | 0.14 | 0.41 | 0.81 | 0.01 | 0.12 | 0.67 | 28.99 | 0.41 | 68.72 | 4.09 | 5.81 | 5.04 | 0.41 | 1.9
| Analytical results for Sr in OREAS 906b (Certified Value 10.8 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 2
| 1 | 1 | 11.27 | 11.4 | 9.8942 | 10.9 | 10.7 | 10.6 | 11 | 11.4 | 11 | 10.2 | 10.1 | 13.6 | 10.6 | 10.6395 | 10.1 | 12 | 8.9 | 10.5 | 10.2 | 11 | 11.3 | 12.4 | 10.8 |  |  |  | 1
| 1 | 2 | 11.32 | 11.6 | 10.0711 | 11.1 | 10.7 | 11.4 | 11 | 11.3 | 10.2 | 10.2 | 10.3 | 14 | 11 | 10.5577 | 9.7 | 11 | 8.1 | 10.7 | 10.4 | 11 | 11 | 12.3 | 10.4 |  |  |  | 25
| 1 | 3 | 11.19 | 11.3 | 10.9645 | 13 | 10.7 | 11.4 | 10.8 | 11.2 | 10.5 | 10.3 | 10.2 | 14 | 13.7 | 10.7103 | 9.5 | 11 | 8.5 | 10.8 | 10.6 | 10 | 11.4 | 12.6 | 10.8 |  |  |  | 16
| 1 | 4 | 11.32 | 11.6 | 10.2392 | 11 | 10.7 | 11 | 10.2 | 11.5 | 10 | 9.9 | 10.8 | 13.6 | 10.8 | 10.755 | 10 | 11 | 8.2 | 10.5 | 10.6 | 10 | 11.6 | 12.7 | 10.6 |  |  |  | 10.7677
| 1 | 5 | 11.1 | 11.6 | 10.8532 | 10.8 | 10.7 | 11.2 | 10.8 | 11.6 | 10.3 | 10.3 | 10.4 | 14 | 10.8 | 10.4633 | 9.8 | 12 | 9.8 | 11 | 10.7 | 10 | 11.8 | 11.8 | 10.6 |  |  |  | 161
| 1 | 6 | 11.41 | 11.4 | 10.3881 | 13.1 | 10.7 | 11.1 | 10.8 | 11.6 | 11 | 10 | 10.5 | 14 | 10.7 | 10.4078 | 10.1 | 11 | 8.5 | 10.7 | 11.3 | 10 | 11 | 12.7 | 10.6
| Mean |  | 11.2683 | 11.4833 | 10.4017 | 11.65 | 10.7 | 11.1167 | 10.7667 | 11.4333 | 10.5 | 10.15 | 10.3833 | 13.8667 | 11.2667 | 10.5889 | 9.8667 | 11.3333 | 8.6667 | 10.7 | 10.6333 | 10.3333 | 11.35 | 12.4167 | 10.6333
| Median |  | 11.295 | 11.5 | 10.3137 | 11.05 | 10.7 | 11.15 | 10.8 | 11.45 | 10.4 | 10.2 | 10.35 | 14 | 10.8 | 10.5986 | 9.9 | 11 | 8.5 | 10.7 | 10.6 | 10 | 11.35 | 12.5 | 10.6
| Std Dev. |  | 0.10944 | 0.13292 | 0.42755 | 1.0895 | 1.9459e-15 | 0.29944 | 0.29439 | 0.1633 | 0.41952 | 0.16432 | 0.24833 | 0.20656 | 1.1994 | 0.13742 | 0.24221 | 0.5164 | 0.62183 | 0.18974 | 0.37238 | 0.5164 | 0.32094 | 0.34303 | 0.15055
| Rel.Std.Dev. |  | 0.9712% | 1.157% | 4.11% | 9.352% | 1.819e-14% | 2.694% | 2.734% | 1.428% | 3.995% | 1.619% | 2.392% | 1.49% | 10.65% | 1.298% | 2.455% | 4.556% | 7.175% | 1.773% | 3.502% | 4.997% | 2.828% | 2.763% | 1.416%
| PDM3 |  | 4.649% | 6.646% | -3.399% | 8.194% | -0.6289% | 3.241% | -0.009797% | 6.182% | -2.486% | -5.737% | -3.57% | 28.78% | 4.634% | -1.66% | -8.368% | 5.253% | -19.51% | -0.6289% | -1.248% | -4.034% | 5.408% | 15.31% | -1.248%
| Z-Score (Absolute) |  | 0.7 | 0.96 | 0.37 | 1.16 | 0 | 0.51 | 0.08 | 0.9 | 0.25 | 0.67 | 0.39 | 3.88 | 0.69 | 0.14 | 1.02 | NA | 2.49 | 0 | 0.08 | NA | 0.8 | 2.1 | 0.08
Results from laboratories 2.19 and 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 906b (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | 0.22 | < 0.01 | < 0.05 | < 0.05 | [10 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | 0.24 | < 0.01 | < 0.05 | < 0.05 | [10 empty] | 36
| 1 | 3 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | 0.39 | < 0.01 | < 0.05 | < 0.05 | [10 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | 0.26 | < 0.01 | < 0.05 | < 0.05 | [10 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | 0.35 | < 0.01 | < 0.05 | < 0.05 | [10 empty] | 162
| 1 | 6 | < 0.05 | < 0.05 | < 5 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | 0.24 | < 0.01 | < 0.05 | < 0.05
| Mean |  |  |  |  | 0.01 | 0.01 | 0.01 | 0.01 | [5 empty] | 0.28333
| Median |  |  |  |  | 0.01 | 0.01 | 0.01 | 0.01 | [5 empty] | 0.25
| Std Dev. | [13 empty] | 0.069474
| Rel.Std.Dev. | [13 empty] | 24.52%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 174.82 | 0.67 | 0.67 | 0.67 | 0.67 | 0.73 | 0.67 | 0.73 | 350.73 | 0.73 | 18.85 | 0.73 | 0.67 | 0.67
| Analytical results for Tb in OREAS 906b (Certified Value 0.43 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.01 | 2.06 | 2.14 | 2.15 | 2.17 | 2.18 | 2.21 | 2.24 | 2.25 | [17 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [17 empty] | 2
|  |  | 01g | 0.25g | 0.25g | 0.2g | 0.2g | 0.5g | 0.25g | 0.25g | 0.5g | [17 empty] | 3
| 1 | 1 | 0.4 | < 5 | 0.44 | 0.41124 | 0.5 | 0.55461 | 0.52 | 0.42 | 0.4 | [17 empty] | 1
| 1 | 2 | 0.398 | < 5 | 0.42 | 0.39992 | 0.4 | 0.51643 | 0.48 | 0.43 | 0.4 | [17 empty] | 37
| 1 | 3 | 0.38 | < 5 | 0.44 | 0.41171 | 0.4 | 0.54188 | 0.51 | 0.43 | 0.4 | [17 empty] | 16
| 1 | 4 | 0.406 | < 5 | 0.43 | 0.41126 | 0.4 | 0.53201 | 0.52 | 0.41 | 0.4 | [17 empty] | 0.43206
| 1 | 5 | 0.374 | < 5 | 0.43 | 0.41541 | 0.4 | 0.52945 | 0.5 | 0.43 | 0.4 | [17 empty] | 163
| 1 | 6 | 0.389 | < 5 | 0.43 | 0.41533 | 0.4 | 0.50975 | 0.5 | 0.41 | 0.4
| Mean |  | 0.39117 |  | 0.43167 | 0.41081 | 0.41667 | 0.53069 | 0.505 | 0.42167 | 0.4
| Median |  | 0.3935 |  | 0.43 | 0.41149 | 0.4 | 0.53073 | 0.505 | 0.425 | 0.4
| Std Dev. |  | 0.0124 |  | 0.0075277 | 0.0056814 | 0.040825 | 0.016385 | 0.015166 | 0.0098319 | 6.0809e-17
| Rel.Std.Dev. |  | 3.17% |  | 1.744% | 1.383% | 9.798% | 3.088% | 3.003% | 2.332% | 1.52e-14%
| PDM3 |  | -9.465% |  | -0.09162% | -4.918% | -3.563% | 22.83% | 16.88% | -2.406% | -7.421%
| Z-Score (Absolute) |  | 0.67 | 34.44 | 0 | 0.35 | NA | 1.65 | 1.22 | 0.17 | NA
Results from laboratories 2.17 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 906b (Certified Value 0.18 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | [6 empty] | 2
| 1 | 1 | 0.2 | 0.14 | 0.14663 | 0.14 | < 5 | 0.17 | 0.19 | 0.18 | 0.18 | 0.17 | 0.18 | < 0.2 | < 10 | 0.17 | 0.06 | < 0.2 | 0.15 | 0.18 | 0.13 | 0.18 | [6 empty] | 1
| 1 | 2 | 0.2 | 0.16 | 0.18554 | 0.22 | < 5 | 0.18 | 0.19 | 0.19 | 0.23 | 0.16 | 0.17 | < 0.2 | < 10 | 0.18 | 0.06 | < 0.2 | 0.17 | 0.2 | 0.15 | 0.2 | [6 empty] | 11
| 1 | 3 | 0.2 | 0.15 | 0.14111 | 0.16 | < 5 | 0.17 | 0.21 | 0.21 | 0.16 | 0.16 | 0.17 | < 0.2 | < 10 | 0.18 | 0.06 | < 0.2 | 0.16 | 0.17 | 0.2 | 0.17 | [6 empty] | 16
| 1 | 4 | 0.2 | 0.17 | 0.19925 | 0.16 | < 5 | 0.14 | 0.16 | 0.2 | 0.18 | 0.16 | 0.18 | < 0.2 | < 10 | 0.19 | 0.05 | < 0.2 | 0.17 | 0.21 | 0.2 | 0.18 | [6 empty] | 0.17518
| 1 | 5 | 0.2 | 0.2 | 0.21145 | 0.17 | < 5 | 0.13 | 0.18 | 0.21 | 0.17 | 0.17 | 0.18 | < 0.2 | < 10 | 0.19 | 0.06 | < 0.2 | 0.15 | 0.22 | 0.12 | 0.19 | [6 empty] | 164
| 1 | 6 | 0.2 | 0.17 | 0.16685 | 0.19 | < 5 | 0.17 | 0.2 | 0.23 | 0.18 | 0.15 | 0.17 | < 0.2 | < 10 | 0.17 | 0.06 | < 0.2 | 0.15 | 0.17 | 0.18 | 0.18
| Mean |  | 0.2 | 0.165 | 0.17514 | 0.17333 |  | 0.16 | 0.18833 | 0.20333 | 0.18333 | 0.16167 | 0.175 |  |  | 0.18 | 0.058333 |  | 0.15833 | 0.19167 | 0.16333 | 0.18333
| Median |  | 0.2 | 0.165 | 0.1762 | 0.165 |  | 0.17 | 0.19 | 0.205 | 0.18 | 0.16 | 0.175 |  |  | 0.18 | 0.06 |  | 0.155 | 0.19 | 0.165 | 0.18
| Std Dev. |  | 3.0405e-17 | 0.020736 | 0.028452 | 0.028048 |  | 0.02 | 0.017224 | 0.017512 | 0.024221 | 0.0075277 | 0.0054772 |  |  | 0.0089443 | 0.0040825 |  | 0.0098319 | 0.02137 | 0.035024 | 0.010328
| Rel.Std.Dev. |  | 1.52e-14% | 12.57% | 16.25% | 16.18% |  | 12.5% | 9.145% | 8.612% | 13.21% | 4.656% | 3.13% |  |  | 4.969% | 6.999% |  | 6.21% | 11.15% | 21.44% | 5.633%
| PDM3 |  | 14.17% | -5.809% | -0.02116% | -1.052% |  | -8.664% | 7.511% | 16.07% | 4.656% | -7.712% | -0.1008% |  |  | 2.753% | -66.7% |  | -9.615% | 9.413% | -6.761% | 4.656%
| Z-Score (Absolute) |  | NA | 0.51 | 0.01 | 0.08 | 117.58 | 0.76 | 0.67 | 1.43 | 0.42 | 0.67 | 0 | 3.79 | 244.02 | 0.25 | 5.9 | 3.79 | 0.84 | 0.84 | 0.59 | 0.42
Results from laboratory 2.01 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 906b (Certified Value 8.32 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 8.38 | 8.83 | 8.6543 | 7.2 | 8.3 | 7.2 | 7.4 | 9.4 | 8.2 | 7.7 | 8 | < 20 | 7.9 | 7.132 | 8.7444 | < 10 | 10.7 | 9.76 | 8.8 | 8.45 | 8.2 | 9.3 | 7.6 |  |  |  | 1
| 1 | 2 | 8.5 | 8.81 | 8.6683 | 7.7 | 8.6 | 7.4 | 7.2 | 9.5 | 8.4 | 8 | 7.95 | < 20 | 7.7 | 7.4578 | 8.6075 | < 10 | 10.3 | 8.8 | 9.2 | 8.875 | 8.4 | 9.3 | 7.8 |  |  |  | 39
| 1 | 3 | 8.16 | 8.8 | 8.9183 | 7.8 | 8.8 | 7.9 | 7.2 | 9 | 8.3 | 8.3 | 8.02 | < 20 | 7.8 | 7.4704 | 8.7192 | < 10 | 10.6 | 9.5 | 9.5 | 8.475 | 8.3 | 9.4 | 7.8 |  |  |  | 16
| 1 | 4 | 8.55 | 8.84 | 8.8562 | 7.2 | 8.3 | 7.5 | 7.8 | 9.4 | 8.3 | 7.8 | 8.24 | < 20 | 7.7 | 7.3393 | 8.6155 | < 10 | 10.9 | 9.12 | 9.1 | 8.91 | 8.2 | 9.2 | 7.7 |  |  |  | 8.3231
| 1 | 5 | 8.08 | 8.9 | 8.8541 | 7.3 | 8.2 | 7.7 | 7.4 | 9.3 | 8.2 | 8.1 | 8.01 | < 20 | 7.7 | 7.7098 | 8.4067 | < 10 | 10.3 | 9.23 | 9.4 | 8.425 | 8.5 | 9.2 | 7.9 |  |  |  | 165
| 1 | 6 | 8.47 | 8.84 | 8.8735 | 7.6 | 8.3 | 7.2 | 7.3 | 9.5 | 8.5 | 7.5 | 8.19 | < 20 | 7.7 | 7.1798 | 8.5462 | < 10 | 10.4 | 9.47 | 8.8 | 8.175 | 8.1 | 9.4 | 7.8
| Mean |  | 8.3567 | 8.8367 | 8.8041 | 7.4667 | 8.4167 | 7.4833 | 7.3833 | 9.35 | 8.3167 | 7.9 | 8.0683 |  | 7.75 | 7.3815 | 8.6066 |  | 10.5333 | 9.3133 | 9.1333 | 8.5517 | 8.2833 | 9.3 | 7.7667
| Median |  | 8.425 | 8.835 | 8.8552 | 7.45 | 8.3 | 7.45 | 7.35 | 9.4 | 8.3 | 7.9 | 8.015 |  | 7.7 | 7.3985 | 8.6115 |  | 10.5 | 9.35 | 9.15 | 8.4625 | 8.25 | 9.3 | 7.8
| Std Dev. |  | 0.19315 | 0.035024 | 0.11311 | 0.26583 | 0.23166 | 0.27869 | 0.22286 | 0.18708 | 0.1169 | 0.28983 | 0.1172 |  | 0.083666 | 0.21271 | 0.12275 |  | 0.24221 | 0.33667 | 0.29439 | 0.28534 | 0.1472 | 0.089443 | 0.10328
| Rel.Std.Dev. |  | 2.311% | 0.3963% | 1.285% | 3.56% | 2.752% | 3.724% | 3.018% | 2.001% | 1.406% | 3.669% | 1.453% |  | 1.08% | 2.882% | 1.426% |  | 2.299% | 3.615% | 3.223% | 3.337% | 1.777% | 0.9617% | 1.33%
| PDM3 |  | 0.4032% | 6.17% | 5.779% | -10.29% | 1.124% | -10.09% | -11.29% | 12.34% | -0.07742% | -5.084% | -3.061% |  | -6.886% | -11.31% | 3.406% |  | 26.56% | 11.9% | 9.735% | 2.746% | -0.4779% | 11.74% | -6.686%
| Z-Score (Absolute) |  | 0 | 0.53 | 0.5 | 0.99 | 0.07 | 0.97 | 1.08 | 1.1 | 0.04 | 0.51 | 0.32 | 1.83 | 0.67 | 1.08 | 0.28 | 3.73 | 2.42 | 1.06 | 0.86 | 0.22 | 0.08 | 1.05 | 0.66
| Analytical results for Ti in OREAS 906b (Certified Value 0.015 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.0125 | 0.0186 | 0.017019 | 0.014 | 0.02 | 0.015 | 0.014 | 0.017 | 0.014 | 0.015 | 0.016 | 0.014 | 0.01 | 0.015777 | 0.017 | 0.01 | < 0.01 | 0.026 | 0.013 | 0.014 | NA | 0.02 | 0.02 | 0.012 |  |  | 1
| 1 | 2 | 0.0124 | 0.019 | 0.017419 | 0.014 | 0.02 | 0.015 | 0.014 | 0.016 | 0.015 | 0.014 | 0.016 | 0.01 | 0.01 | 0.015788 | 0.018 | 0.01 | < 0.01 | 0.0231 | 0.013 | 0.014 | NA | 0.02 | 0.02 | 0.012 |  |  | 40
| 1 | 3 | 0.0123 | 0.0196 | 0.018101 | 0.016 | 0.02 | 0.016 | 0.014 | 0.016 | 0.014 | 0.014 | 0.017 | 0.01 | 0.01 | 0.016671 | 0.018 | 0.01 | < 0.01 | 0.0234 | 0.013 | 0.014 | NA | 0.02 | 0.02 | 0.012 |  |  | 16
| 1 | 4 | 0.0128 | 0.0191 | 0.017657 | 0.014 | 0.02 | 0.015 | 0.014 | 0.016 | 0.014 | 0.014 | 0.017 | 0.012 | 0.01 | 0.016282 | 0.017 | 0.02 | < 0.01 | 0.0238 | 0.014 | 0.015 | NA | 0.02 | 0.02 | 0.013 |  |  | 0.015324
| 1 | 5 | 0.012 | 0.0194 | 0.017018 | 0.014 | 0.02 | 0.016 | 0.014 | 0.017 | 0.014 | 0.014 | 0.016 | 0.01 | 0.01 | 0.01643 | 0.018 | 0.02 | < 0.01 | 0.026 | 0.014 | 0.015 | NA | 0.02 | 0.02 | 0.013 |  |  | 166
| 1 | 6 | 0.0125 | 0.0186 | 0.018183 | 0.017 | 0.02 | 0.015 | 0.013 | 0.016 | 0.014 | 0.015 | 0.017 | 0.01 | 0.01 | 0.015916 | 0.017 | 0.01 | < 0.01 | 0.0249 | 0.013 | 0.016 | NA | 0.02 | 0.02 | 0.013
| Mean |  | 0.012417 | 0.01905 | 0.017566 | 0.014833 | 0.02 | 0.015333 | 0.013833 | 0.016333 | 0.014167 | 0.014333 | 0.0165 | 0.011 | 0.01 | 0.016144 | 0.0175 | 0.013333 |  | 0.024533 | 0.013333 | 0.014667 |  | 0.02 | 0.02 | 0.0125
| Median |  | 0.01245 | 0.01905 | 0.017538 | 0.014 | 0.02 | 0.015 | 0.014 | 0.016 | 0.014 | 0.014 | 0.0165 | 0.01 | 0.01 | 0.016099 | 0.0175 | 0.01 |  | 0.02435 | 0.013 | 0.0145 |  | 0.02 | 0.02 | 0.0125
| Std Dev. |  | 0.00026394 | 0.00040866 | 0.00050944 | 0.0013292 | 0 | 0.0005164 | 0.00040825 | 0.0005164 | 0.00040825 | 0.0005164 | 0.00054772 | 0.0016733 | 0 | 0.00037197 | 0.00054772 | 0.005164 |  | 0.0012894 | 0.0005164 | 0.0008165 |  | 0 | 0 | 0.00054772
| Rel.Std.Dev. |  | 2.126% | 2.145% | 2.9% | 8.961% | 0% | 3.368% | 2.951% | 3.162% | 2.882% | 3.603% | 3.32% | 15.21% | 0% | 2.304% | 3.13% | 38.73% |  | 5.256% | 3.873% | 5.567% |  | 0% | 0% | 4.382%
| PDM3 |  | -18.97% | 24.31% | 14.63% | -3.204% | 30.51% | 0.0591% | -9.729% | 6.585% | -7.554% | -6.466% | 7.672% | -28.22% | -34.74% | 5.348% | 14.2% | -12.99% |  | 60.09% | -12.99% | -4.291% |  | 30.51% | 30.51% | -18.43%
| Z-Score (Absolute) |  | 0.7 | 1.22 | 0.79 | 0 | 1.49 | 0.14 | 0.29 | 0.43 | 0.19 | 0.14 | 0.48 | 1.11 | 1.4 | 0.38 | 0.77 | 0.43 | 2.84 | 2.8 | 0.43 | 0.05 | NA | 1.49 | 1.49 | 0.67
| Analytical results for Tl in OREAS 906b (Certified Value 0.073 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.07 | 0.09 | 0.072296 | 0.06 | < 5 | 0.06 | 0.07 | 0.08 | 0.06 | 0.07 | 0.08 | < 10 | 0.07 | < 0.1 | 0.12 | < 5 | 0.16 | 0.07 | 0.04 | 0.1 | 0.07 | 0.09 | 0.068 | 0.07 |  |  | 1
| 1 | 2 | 0.07 | 0.09 | 0.075828 | 0.07 | < 5 | 0.06 | 0.07 | 0.07 | 0.06 | 0.07 | 0.08 | < 10 | 0.07 | < 0.1 | 0.12 | < 5 | 0.14 | 0.07 | 0.04 | 0.1 | 0.07 | 0.09 | 0.063 | 0.07 |  |  | 41
| 1 | 3 | 0.07 | 0.09 | 0.080733 | 0.08 | < 5 | 0.07 | 0.07 | 0.07 | 0.06 | 0.06 | 0.08 | < 10 | 0.07 | < 0.1 | 0.12 | < 5 | 0.14 | 0.08 | 0.04 | 0.1 | 0.07 | 0.09 | 0.066 | 0.07 |  |  | 16
| 1 | 4 | 0.08 | 0.09 | 0.075645 | 0.06 | < 5 | 0.07 | 0.07 | 0.08 | 0.06 | 0.06 | 0.08 | < 10 | 0.06 | < 0.1 | 0.1 | < 5 | 0.16 | 0.09 | 0.05 | 0.1 | 0.07 | 0.09 | 0.068 | 0.07 |  |  | 0.072655
| 1 | 5 | 0.07 | 0.09 | 0.079422 | 0.06 | < 5 | 0.07 | 0.07 | 0.08 | 0.07 | 0.06 | 0.08 | < 10 | 0.07 | < 0.1 | 0.1 | < 5 | 0.13 | 0.07 | 0.04 | 0.1 | 0.07 | 0.09 | 0.068 | 0.07 |  |  | 167
| 1 | 6 | 0.07 | 0.09 | 0.074976 | 0.07 | < 5 | 0.07 | 0.06 | 0.08 | 0.06 | 0.07 | 0.08 | < 10 | 0.06 | < 0.1 | 0.1 | < 5 | 0.16 | 0.08 | 0.05 | 0.1 | 0.07 | 0.09 | 0.063 | 0.07
| Mean |  | 0.071667 | 0.09 | 0.076483 | 0.066667 |  | 0.066667 | 0.068333 | 0.076667 | 0.061667 | 0.065 | 0.08 |  | 0.066667 |  | 0.11 |  | 0.14833 | 0.076667 | 0.043333 | 0.1 | 0.07 | 0.09 | 0.066 | 0.07
| Median |  | 0.07 | 0.09 | 0.075736 | 0.065 |  | 0.07 | 0.07 | 0.08 | 0.06 | 0.065 | 0.08 |  | 0.07 |  | 0.11 |  | 0.15 | 0.075 | 0.04 | 0.1 | 0.07 | 0.09 | 0.067 | 0.07
| Std Dev. |  | 0.0040825 | 1.5202e-17 | 0.0030863 | 0.008165 |  | 0.005164 | 0.0040825 | 0.005164 | 0.0040825 | 0.0054772 | 0 |  | 0.005164 |  | 0.010954 |  | 0.013292 | 0.008165 | 0.005164 | 1.5202e-17 | 0 | 1.5202e-17 | 0.0024495 | 0
| Rel.Std.Dev. |  | 5.696% | 1.689e-14% | 4.035% | 12.25% |  | 7.746% | 5.974% | 6.736% | 6.62% | 8.427% | 0% |  | 7.746% |  | 9.959% |  | 8.961% | 10.65% | 11.92% | 1.52e-14% | 0% | 1.689e-14% | 3.711% | 0%
| PDM3 |  | -1.361% | 23.87% | 5.269% | -8.242% |  | -8.242% | -5.948% | 5.521% | -15.12% | -10.54% | 10.11% |  | -8.242% |  | 51.4% |  | 104.2% | 5.521% | -40.36% | 37.64% | -3.655% | 23.87% | -9.16% | -3.655%
| Z-Score (Absolute) |  | 0 | 1.85 | 0.49 | 0.51 | 245.62 | 0.51 | 0.34 | 0.51 | 1.01 | 0.67 | 0.84 | 498.48 | 0.51 | 2.19 | 3.88 | 245.62 | 7.75 | 0.51 | 2.87 | NA | 0.17 | 1.85 | 0.57 | 0.17
Results from laboratory 2.23 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Tm in OREAS 906b (Indicative Value 0.043 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | [23 empty] | 1
| No | No | 2.01 | 2.15 | 2.25 | [23 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 3
|  |  | 01g | 0.2g | 0.5g | [23 empty] | 3
| 1 | 1 | 0.043 | < 0.05 | < 0.1 | [23 empty] | 1
| 1 | 2 | 0.042 | < 0.05 | < 0.1 | [23 empty] | 19
| 1 | 3 | 0.043 | < 0.05 | < 0.1 | [23 empty] | 16
| 1 | 4 | 0.044 | < 0.05 | < 0.1 | [23 empty] | 0.042833
| 1 | 5 | 0.04 | < 0.05 | < 0.1 | [23 empty] | 25
| 1 | 6 | 0.045 | < 0.05 | < 0.1
| Mean |  | 0.042833
| Median |  | 0.043
| Std Dev. |  | 0.0017224
| Rel.Std.Dev. |  | 4.021%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | 0 | 1.68 | 0.67
| Analytical results for U in OREAS 906b (Certified Value 2.23 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 2.2 | 2.55 | 2.3735 | 1.88 | < 5 | 1.95 | 2.1 | 2.39 | 2 | 2.05 | 2.45 | < 10 | 2.13 | 2.1774 | 2.14 | 2.4326 | < 10 | 2.96 | 2.53 | 2.23 | 2.13 | 2.23 | 2.5 | 1.84 | 2.1 |  | 1
| 1 | 2 | 2.24 | 2.55 | 2.4106 | 2.04 | < 5 | 2.04 | 2.07 | 2.36 | 2.02 | 2.09 | 2.43 | < 10 | 2.01 | 2.2481 | 2.07 | 2.408 | < 10 | 2.94 | 2.5 | 2.31 | 2.15 | 2.28 | 2.5 | 1.78 | 2.2 |  | 43
| 1 | 3 | 2.18 | 2.57 | 2.7124 | 2.14 | < 5 | 2.17 | 2.09 | 2.32 | 1.97 | 2.09 | 2.46 | < 10 | 2.1 | 2.346 | 2.05 | 2.4338 | < 10 | 2.94 | 2.56 | 2.37 | 2.31 | 2.29 | 2.5 | 1.75 | 2.1 |  | 16
| 1 | 4 | 2.24 | 2.61 | 2.4545 | 1.95 | < 5 | 2.03 | 2.04 | 2.34 | 2 | 2.05 | 2.54 | < 10 | 2.01 | 2.284 | 2.17 | 2.4187 | < 10 | 2.98 | 2.49 | 2.2 | 2.04 | 2.21 | 2.4 | 1.88 | 2.1 |  | 2.2312
| 1 | 5 | 2.11 | 2.63 | 2.6719 | 2.02 | < 5 | 2.08 | 2.12 | 2.38 | 1.96 | 2.07 | 2.54 | < 10 | 2.04 | 2.4339 | 2.1 | 2.3661 | < 10 | 2.86 | 2.44 | 2.3 | 2.21 | 2.29 | 2.5 | 1.95 | 2.1 |  | 168
| 1 | 6 | 2.22 | 2.62 | 2.4254 | 2.1 | < 5 | 1.89 | 2.11 | 2.37 | 2.01 | 2 | 2.57 | < 10 | 2.06 | 2.2982 | 2.17 | 2.2931 | < 10 | 2.9 | 2.54 | 2.23 | 2.12 | 2.24 | 2.5 | 1.77 | 2.1
| Mean |  | 2.1983 | 2.5883 | 2.508 | 2.0217 |  | 2.0267 | 2.0883 | 2.36 | 1.9933 | 2.0583 | 2.4983 |  | 2.0583 | 2.298 | 2.1167 | 2.392 |  | 2.93 | 2.51 | 2.2733 | 2.16 | 2.2567 | 2.4833 | 1.8283 | 2.1167
| Median |  | 2.21 | 2.59 | 2.4399 | 2.03 |  | 2.035 | 2.095 | 2.365 | 2 | 2.06 | 2.5 |  | 2.05 | 2.2911 | 2.12 | 2.4134 |  | 2.94 | 2.515 | 2.265 | 2.14 | 2.26 | 2.5 | 1.81 | 2.1
| Std Dev. |  | 0.04916 | 0.036009 | 0.14553 | 0.095586 |  | 0.098116 | 0.029269 | 0.026077 | 0.023381 | 0.033714 | 0.058452 |  | 0.048751 | 0.087188 | 0.051251 | 0.054443 |  | 0.043359 | 0.042895 | 0.064083 | 0.091652 | 0.034448 | 0.040825 | 0.07679 | 0.040825
| Rel.Std.Dev. |  | 2.236% | 1.391% | 5.803% | 4.728% |  | 4.841% | 1.402% | 1.105% | 1.173% | 1.638% | 2.34% |  | 2.368% | 3.794% | 2.421% | 2.276% |  | 1.48% | 1.709% | 2.819% | 4.243% | 1.527% | 1.644% | 4.2% | 1.929%
| PDM3 |  | -1.472% | 16.01% | 12.41% | -9.39% |  | -9.166% | -6.402% | 5.774% | -10.66% | -7.746% | 11.97% |  | -7.746% | 2.994% | -5.132% | 7.21% |  | 31.32% | 12.5% | 1.89% | -3.19% | 1.143% | 11.3% | -18.05% | -5.132%
| Z-Score (Absolute) |  | 0.24 | 0.99 | 0.74 | 0.79 | 0.71 | 0.77 | 0.58 | 0.27 | 0.88 | 0.67 | 0.71 | 8.55 | 0.67 | 0.08 | 0.49 | 0.37 | 8.55 | 2.06 | 0.74 | 0 | 0.36 | 0.05 | 0.66 | 1.4 | 0.49
| Analytical results for V in OREAS 906b (Certified Value 2.01 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  |  | 3
| 1 | 1 | 4 | 2.0011 | 2 | 5.8 | 2 | 2 | 2 | 2 | 2 | 2 | 2.3 | 4 | 1.8531 | 2 | 3 | 2 | 1 | 2 | < 1 | 2 | 2 | 2 |  |  |  |  | 1
| 1 | 2 | 4 | 2.1378 | 2 | 5.9 | 2 | 2 | 2 | 2 | 1 | 2 | 2 | 4 | 1.9663 | 2 | 3 | 2 | 3 | 2 | < 1 | 2 | 2 | 2 |  |  |  |  | 12
| 1 | 3 | 4 | 2.2998 | 2 | 6 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 1.9467 | 2 | 3 | 2 | 3 | 2 | < 1 | 3 | 2 | 2 |  |  |  |  | 16
| 1 | 4 | 4 | 2.0508 | 2 | 5.9 | 2 | 2 | 2 | 2 | 2 | 2 | 2.3 | 4 | 1.9823 | 3 | 2.5 | 2 | 1 | 2 | < 1 | 2 | 2 | 2 |  |  |  |  | 2.0127
| 1 | 5 | 4 | 2.1083 | 2 | 6 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 1.9891 | 2 | 3 | 2 | 2 | 2 | 1 | 3 | 2 | 2 |  |  |  |  | 169
| 1 | 6 | 4 | 2.3656 | 2 | 5.7 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 4 | 1.9134 | 3 | 3 | 2 | 3 | 2 | < 1 | 3 | 2 | 2
| Mean |  | 4 | 2.1606 | 2 | 5.8833 | 2 | 2 | 2 | 2 | 1.8333 | 2 | 2.1 | 3.6667 | 1.9418 | 2.3333 | 2.9167 | 2 | 2.1667 | 2 | 1 | 2.5 | 2 | 2
| Median |  | 4 | 2.1231 | 2 | 5.9 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 4 | 1.9565 | 2 | 3 | 2 | 2.5 | 2 | 1 | 2.5 | 2 | 2
| Std Dev. |  | 0 | 0.14294 | 0 | 0.1169 | 0 | 0 | 0 | 0 | 0.40825 | 0 | 0.15492 | 0.5164 | 0.05133 | 0.5164 | 0.20412 | 0 | 0.98319 | 0 |  | 0.54772 | 0 | 0
| Rel.Std.Dev. |  | 0% | 6.616% | 0% | 1.987% | 0% | 0% | 0% | 0% | 22.27% | 0% | 7.377% | 14.08% | 2.643% | 22.13% | 6.999% | 0% | 45.38% | 0% |  | 21.91% | 0% | 0%
| PDM3 |  | 98.74% | 7.35% | -0.6286% | 192.3% | -0.6286% | -0.6286% | -0.6286% | -0.6286% | -8.91% | -0.6286% | 4.34% | 82.18% | -3.519% | 15.93% | 44.92% | -0.6286% | 7.652% | -0.6286% | -50.31% | 24.21% | -0.6286% | -0.6286%
| Z-Score (Absolute) |  | 17.05 | 1.37 | 0 | 33.11 | 0 | 0 | 0 | 0 | 1.42 | 0 | 0.85 | 14.21 | 0.5 | 2.84 | 7.82 | 0 | 1.42 | 0 | 12.08 | 4.26 | 0 | 0
| Analytical results for W in OREAS 906b (Certified Value 0.66 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 2
| 1 | 1 | 0.63 | 0.7 | 0.76492 | 0.49 | < 5 | 0.65 | 0.76 | 0.71 | 0.56 | 0.6 | 0.46 | < 10 | 0.7 | < 1 | < 10 | 0.7 | < 20 | 0.98 | 1.81 | 0.72 | 0.5 | 0.7 | 0.6 |  |  |  | 1
| 1 | 2 | 0.6 | 0.7 | 0.6764 | 0.56 | < 5 | 0.65 | 0.72 | 0.71 | 0.6 | 0.62 | 0.45 | < 10 | 0.7 | < 1 | < 10 | 0.7 | < 20 | 0.84 | 1.61 | 0.81 | 0.5 | 0.5 | 0.7 |  |  |  | 13
| 1 | 3 | 0.56 | 0.7 | 0.81455 | 0.64 | < 5 | 0.69 | 0.69 | 0.7 | 0.59 | 0.58 | 0.43 | < 10 | 0.8 | < 1 | < 10 | 0.8 | < 20 | 0.84 | 1.54 | 0.81 | 0.5 | 0.6 | 0.6 |  |  |  | 16
| 1 | 4 | 0.61 | 0.8 | 0.67202 | 0.52 | < 5 | 0.65 | 0.72 | 0.68 | 0.56 | 0.59 | 0.45 | < 10 | 0.7 | < 1 | < 10 | 0.6 | < 20 | 0.96 | 1.84 | 0.81 | 0.5 | 0.6 | 0.6 |  |  |  | 0.66228
| 1 | 5 | 0.59 | 0.7 | 0.85309 | 0.55 | < 5 | 0.66 | 0.75 | 0.71 | 0.57 | 0.7 | 0.48 | < 10 | 0.7 | < 1 | < 10 | 0.6 | < 20 | 0.89 | 1.63 | 0.69 | 0.5 | 0.6 | 0.7 |  |  |  | 170
| 1 | 6 | 0.59 | 0.7 | 0.67751 | 0.53 | < 5 | 0.71 | 0.71 | 0.74 | 0.66 | 0.62 | 0.42 | < 10 | 0.7 | < 1 | < 10 | 0.7 | < 20 | 0.86 | 1.89 | 0.72 | 0.6 | 0.7 | 0.7
| Mean |  | 0.59667 | 0.71667 | 0.74308 | 0.54833 |  | 0.66833 | 0.725 | 0.70833 | 0.59 | 0.61833 | 0.44833 |  | 0.71667 |  |  | 0.68333 |  | 0.895 | 1.72 | 0.76 | 0.51667 | 0.61667 | 0.65
| Median |  | 0.595 | 0.7 | 0.72121 | 0.54 |  | 0.655 | 0.72 | 0.71 | 0.58 | 0.61 | 0.45 |  | 0.7 |  |  | 0.7 |  | 0.875 | 1.72 | 0.765 | 0.5 | 0.6 | 0.65
| Std Dev. |  | 0.023381 | 0.040825 | 0.079348 | 0.051153 |  | 0.025626 | 0.025884 | 0.019408 | 0.037947 | 0.043089 | 0.02137 |  | 0.040825 |  |  | 0.075277 |  | 0.061237 | 0.14422 | 0.055857 | 0.040825 | 0.075277 | 0.054772
| Rel.Std.Dev. |  | 3.919% | 5.696% | 10.68% | 9.329% |  | 3.834% | 3.57% | 2.74% | 6.432% | 6.969% | 4.766% |  | 5.696% |  |  | 11.02% |  | 6.842% | 8.385% | 7.35% | 7.902% | 12.21% | 8.427%
| PDM3 |  | -9.907% | 8.212% | 12.2% | -17.21% |  | 0.914% | 9.47% | 6.954% | -10.91% | -6.636% | -32.3% |  | 8.212% |  |  | 3.179% |  | 35.14% | 159.7% | 14.76% | -21.99% | -6.887% | -1.854%
| Z-Score (Absolute) |  | 0.51 | NA | 0.07 | 0.7 | 7.04 | 0.22 | 0 | 0.07 | 0.54 | 0.42 | 1.1 | 16.96 | NA | 0.89 | 16.96 | NA | 36.79 | 0.67 | 3.95 | 0.14 | NA | NA | NA
Results from laboratories 2.02, 2.14, 2.17, 2.24, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 906b (Certified Value 6.7 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [5 empty] | 2
| 1 | 1 | 6.38 | 7.38 | 6.596 | 5.7 | 5.79 | 5.94 | 6.97 | 5.66 | 5.93 | 6.8 | 5.92 | 6.6185 | 7.8 | 6.2 | 8.3078 | 6 | 8.13 | 7.98 | 6.03 | 6.3 | 7 | [5 empty] | 1
| 1 | 2 | 6.27 | 7.2 | 6.6492 | 5.5 | 5.98 | 6.03 | 6.67 | 5.89 | 5.93 | 6.84 | 6.2 | 6.5506 | 8 | 6 | 8.3492 | 6 | 7.89 | 7.08 | 6.06 | 6.47 | 7.04 | [5 empty] | 46
| 1 | 3 | 6.21 | 7.22 | 7.0309 | 5.8 | 6.49 | 6 | 6.84 | 5.78 | 6.07 | 6.95 | 5.99 | 6.6867 | 7.9 | 5.9 | 8.5869 | 5 | 8.24 | 7.72 | 6.31 | 6.43 | 7.05 | [5 empty] | 16
| 1 | 4 | 6.52 | 7.49 | 6.7938 | 5.7 | 5.84 | 6.04 | 6.78 | 5.68 | 5.8 | 7.17 | 5.97 | 6.6765 | 8.4 | 6.3 | 8.1476 | 6 | 7.77 | 7.86 | 6.25 | 6.12 | 6.84 | [5 empty] | 6.6993
| 1 | 5 | 6.07 | 7.34 | 7.1317 | 5.8 | 5.94 | 6.06 | 7.17 | 5.63 | 6.02 | 6.98 | 6.01 | 6.703 | 8.2 | 6.3 | 8.3772 | 6 | 8.08 | 7.54 | 6.44 | 6.36 | 6.94 | [5 empty] | 171
| 1 | 6 | 6.37 | 7.27 | 6.807 | 5.8 | 5.8 | 5.98 | 6.96 | 5.85 | 5.87 | 7.06 | 5.91 | 6.5327 | 7.6 | 6.3 | 8.1502 | 5 | 8.37 | 7.61 | 6.49 | 6.37 | 7.2
| Mean |  | 6.3033 | 7.3167 | 6.8348 | 5.7167 | 5.9733 | 6.0083 | 6.8983 | 5.7483 | 5.9367 | 6.9667 | 6 | 6.628 | 7.9833 | 6.1667 | 8.3198 | 5.6667 | 8.08 | 7.6317 | 6.2633 | 6.3417 | 7.0117
| Median |  | 6.32 | 7.305 | 6.8004 | 5.75 | 5.89 | 6.015 | 6.9 | 5.73 | 5.93 | 6.965 | 5.98 | 6.6475 | 7.95 | 6.25 | 8.3285 | 6 | 8.105 | 7.665 | 6.28 | 6.365 | 7.02
| Std Dev. |  | 0.15591 | 0.1093 | 0.21001 | 0.1169 | 0.2644 | 0.044008 | 0.17452 | 0.10759 | 0.097912 | 0.13736 | 0.10545 | 0.072917 | 0.28577 | 0.17512 | 0.16375 | 0.5164 | 0.22109 | 0.31448 | 0.19012 | 0.12352 | 0.1204
| Rel.Std.Dev. |  | 2.473% | 1.494% | 3.073% | 2.045% | 4.426% | 0.7324% | 2.53% | 1.872% | 1.649% | 1.972% | 1.758% | 1.1% | 3.58% | 2.84% | 1.968% | 9.113% | 2.736% | 4.121% | 3.035% | 1.948% | 1.717%
| PDM3 |  | -5.91% | 9.215% | 2.022% | -14.67% | -10.84% | -10.31% | 2.971% | -14.19% | -11.38% | 3.991% | -10.44% | -1.064% | 19.17% | -7.951% | 24.19% | -15.41% | 20.61% | 13.92% | -6.508% | -5.338% | 4.663%
| Z-Score (Absolute) |  | 0.25 | 1.13 | 0.47 | 1.04 | 0.69 | 0.64 | 0.56 | 1 | 0.74 | 0.65 | 0.66 | 0.19 | 2.03 | 0.43 | 2.48 | NA | 2.16 | 1.55 | 0.3 | 0.19 | 0.71
| Analytical results for Yb in OREAS 906b (Certified Value 0.23 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.01 | 2.06 | 2.14 | 2.15 | 2.17 | 2.18 | 2.21 | 2.24 | 2.25 | [17 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [17 empty] | 2
|  |  | 01g | 0.25g | 0.25g | 0.2g | 0.2g | 0.5g | 0.25g | 0.25g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.23 | < 5 | 0.2 | 0.2175 | 0.2 | 0.30824 | 0.3 | 0.3 | 0.2 | [17 empty] | 1
| 1 | 2 | 0.239 | < 5 | 0.2 | 0.20623 | 0.2 | 0.29043 | 0.2 | 0.3 | 0.2 | [17 empty] | 47
| 1 | 3 | 0.223 | < 5 | 0.2 | 0.21556 | 0.2 | 0.2877 | 0.2 | 0.3 | 0.2 | [17 empty] | 16
| 1 | 4 | 0.239 | < 5 | 0.2 | 0.21733 | 0.2 | 0.29694 | 0.3 | 0.2 | 0.2 | [17 empty] | 0.23387
| 1 | 5 | 0.22 | < 5 | 0.2 | 0.21629 | 0.2 | 0.30206 | 0.2 | 0.3 | 0.2 | [17 empty] | 172
| 1 | 6 | 0.225 | < 5 | 0.2 | 0.22172 | 0.2 | 0.2583 | 0.3 | 0.3 | 0.2
| Mean |  | 0.22933 |  | 0.2 | 0.21577 | 0.2 | 0.29061 | 0.25 | 0.28333 | 0.2
| Median |  | 0.2275 |  | 0.2 | 0.21681 | 0.2 | 0.29368 | 0.25 | 0.3 | 0.2
| Std Dev. |  | 0.008165 |  | 3.0405e-17 | 0.0051403 | 3.0405e-17 | 0.017518 | 0.054772 | 0.040825 | 3.0405e-17
| Rel.Std.Dev. |  | 3.56% |  | 1.52e-14% | 2.382% | 1.52e-14% | 6.028% | 21.91% | 14.41% | 1.52e-14%
| PDM3 |  | -1.94% |  | -14.48% | -7.738% | -14.48% | 24.26% | 6.897% | 21.15% | -14.48%
| Z-Score (Absolute) |  | 0 | 52.2 | 0.67 | 0.31 | 0.67 | 1.41 | 0.48 | 1.24 | 0.67
| Analytical results for Zn in OREAS 906b (Certified Value 68 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 1
| 1 | 1 | 63 | 80 | 68.4088 | 63 | 60.1 | 67 | 65 | 73 | 65 | 67 | 69 | 107 | 67 | 71.1668 | 67 | 67.7 | 69.24 | 65 | 79.4 | 66 | 88.2 | 69 | 73.6 | 70 | 69.4 |  | 1
| 1 | 2 | 63 | 79 | 67.9607 | 64.2 | 59.2 | 68 | 65 | 71 | 66 | 65 | 71 | 87 | 69 | 71.1821 | 69 | 65.4 | 67.69 | 64 | 83.9 | 70 | 87.2 | 67 | 75.5 | 69 | 69.1 |  | 28
| 1 | 3 | 62 | 79 | 69.6874 | 69.6 | 60.8 | 69 | 66 | 70 | 66 | 65 | 70 | 85 | 69 | 72.0943 | 66 | 66 | 69.48 | 64 | 84.3 | 66 | 89.1 | 66 | 72.5 | 68 | 67.5 |  | 16
| 1 | 4 | 63 | 80 | 69.2967 | 66.5 | 59.1 | 69 | 62 | 72 | 67 | 65 | 72 | 93 | 70 | 70.8504 | 68 | 67.8 | 67.81 | 65 | 79 | 70 | 90 | 67 | 72.2 | 70 | 69.4 |  | 67.5471
| 1 | 5 | 61 | 80 | 70.6109 | 64.8 | 60.8 | 69 | 65 | 72 | 66 | 64 | 70 | 84 | 71 | 72.6211 | 70 | 66.6 | 69 | 65 | 79.9 | 67 | 88.9 | 68 | 77.5 | 71 | 68.7 |  | 173
| 1 | 6 | 62 | 80 | 70.0706 | 68 | 61.1 | 69 | 63 | 71 | 66 | 65 | 72 | 106 | 66 | 70.0699 | 67 | 67 | 66.8 | 63 | 83.5 | 69 | 90 | 66 | 72.3 | 68 | 68.5
| Mean |  | 62.3333 | 79.6667 | 69.3392 | 66.0167 | 60.1833 | 68.5 | 64.3333 | 71.5 | 66 | 65.1667 | 70.6667 | 93.6667 | 68.6667 | 71.3308 | 67.8333 | 66.75 | 68.3367 | 64.3333 | 81.6667 | 68 | 88.9 | 67.1667 | 73.9333 | 69.3333 | 68.7667
| Median |  | 62.5 | 80 | 69.4921 | 65.65 | 60.45 | 69 | 65 | 71.5 | 66 | 65 | 70.5 | 90 | 69 | 71.1745 | 67.5 | 66.8 | 68.405 | 64.5 | 81.7 | 68 | 89 | 67 | 73.05 | 69.5 | 68.9
| Std Dev. |  | 0.8165 | 0.5164 | 1.0041 | 2.4855 | 0.86583 | 0.83666 | 1.5055 | 1.0488 | 0.63246 | 0.98319 | 1.2111 | 10.4243 | 1.8619 | 0.90746 | 1.472 | 0.94604 | 1.0602 | 0.8165 | 2.476 | 1.8974 | 1.0807 | 1.169 | 2.1454 | 1.2111 | 0.72019
| Rel.Std.Dev. |  | 1.31% | 0.6482% | 1.448% | 3.765% | 1.439% | 1.221% | 2.34% | 1.467% | 0.9583% | 1.509% | 1.714% | 11.13% | 2.712% | 1.272% | 2.17% | 1.417% | 1.551% | 1.269% | 3.032% | 2.79% | 1.216% | 1.741% | 2.902% | 1.747% | 1.047%
| PDM3 |  | -7.719% | 17.94% | 2.653% | -2.266% | -10.9% | 1.411% | -4.758% | 5.852% | -2.29% | -3.524% | 4.618% | 38.67% | 1.657% | 5.601% | 0.4237% | -1.18% | 1.169% | -4.758% | 20.9% | 0.6704% | 31.61% | -0.5633% | 9.454% | 2.644% | 1.805%
| Z-Score (Absolute) |  | 1.66 | 3.01 | 0.23 | 0.67 | 2.24 | 0 | 1.12 | 0.81 | 0.67 | 0.9 | 0.58 | 6.79 | 0.04 | 0.76 | 0.18 | 0.47 | 0.04 | 1.12 | 3.55 | 0.13 | 5.5 | 0.36 | 1.47 | 0.22 | 0.07
| Analytical results for Zr in OREAS 906b (Certified Value 46.4 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [7 empty] | 1
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [7 empty] | 2
| 1 | 1 | 28.9 | 52.9 | 48.3147 | 54 | 41.4 | 42.7 | 44.9 | 40.5 | 44.7 | 45.6 | 38.4 | 47.6468 | 45.9 | 30 | 65.7 | 49.2 | 46.6 | 47.2 | 3.1 | [7 empty] | 1
| 1 | 2 | 28.5 | 52 | 49.8554 | 52.9 | 41.6 | 43.8 | 45.5 | 44.6 | 43.6 | 45.1 | 39.4 | 46.8001 | 46.7 | 30 | 64.9 | 46 | 47.9 | 47.7 | 2.3 | [7 empty] | 14
| 1 | 3 | 24.3 | 51.1 | 52.6793 | 54.5 | 42.6 | 43 | 45.6 | 42.5 | 42.8 | 46.1 | 37.9 | 47.8236 | 46.6 | 28 | 61.3 | 52.1 | 48.6 | 48.3 | 3.1 | [7 empty] | 16
| 1 | 4 | 30.3 | 54 | 49.8549 | 54 | 40.8 | 43 | 43.4 | 41.7 | 41.7 | 48 | 38.8 | 47.7779 | 46.8 | 29 | 64.8 | 47.8 | 48.8 | 48.4 | 2.8 | [7 empty] | 46.3833
| 1 | 5 | 26.4 | 53.5 | 52.8517 | 55.3 | 42.6 | 42.2 | 46 | 40.3 | 43.6 | 46.7 | 39.5 | 47.8105 | 46.5 | 30 | 61.8 | 53.3 | 49.7 | 47.3 | 4.5 | [7 empty] | 174
| 1 | 6 | 29.2 | 52 | 50.219 | 53.4 | 42.5 | 42.6 | 46 | 41.4 | 43.3 | 48 | 38.2 | 47.061 | 45.6 | 28 | 66.1 | 40.7 | 46.2 | 47.5 | 3.6
| Mean |  | 27.9333 | 52.5833 | 50.6292 | 54.0167 | 41.9167 | 42.8833 | 45.2333 | 41.8333 | 43.2833 | 46.5833 | 38.7 | 47.4866 | 46.35 | 29.1667 | 64.1 | 48.1833 | 47.9667 | 47.7333 | 3.2333
| Median |  | 28.7 | 52.45 | 50.0372 | 54 | 42.05 | 42.85 | 45.55 | 41.55 | 43.45 | 46.4 | 38.6 | 47.7123 | 46.55 | 29.5 | 64.85 | 48.5 | 48.25 | 47.6 | 3.1
| Std Dev. |  | 2.1915 | 1.0797 | 1.7814 | 0.83766 | 0.76004 | 0.53821 | 0.98522 | 1.5769 | 0.99482 | 1.2189 | 0.65115 | 0.44303 | 0.48477 | 0.98319 | 2.0406 | 4.549 | 1.3486 | 0.50859 | 0.75277
| Rel.Std.Dev. |  | 7.845% | 2.053% | 3.518% | 1.551% | 1.813% | 1.255% | 2.178% | 3.77% | 2.298% | 2.617% | 1.683% | 0.933% | 1.046% | 3.371% | 3.183% | 9.441% | 2.811% | 1.065% | 23.28%
| PDM3 |  | -39.78% | 13.37% | 9.154% | 16.46% | -9.63% | -7.546% | -2.479% | -9.809% | -6.683% | 0.4313% | -16.56% | 2.379% | -0.07174% | -37.12% | 38.2% | 3.881% | 3.414% | 2.911% | -93.03%
| Z-Score (Absolute) |  | 2.9 | 0.98 | 0.67 | 1.21 | 0.7 | 0.55 | 0.18 | 0.71 | 0.48 | 0.04 | 1.21 | 0.18 | 0 | 2.71 | 2.8 | 0.29 | 0.25 | 0.22 | 6.79
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1171: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 906b (Indicative Value 14.58 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.58 | 1
| 1 | 2 | 14.57 | 22
| Mean |  | 14.575 | 16
| Median |  | 14.575 | 14.575
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.04852%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 906b (Indicative Value 10 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 10
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 906b (Indicative Value 3495 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 3495 | 1
| 1 | 2 | 3495 | 24
| Mean |  | 3495 | 16
| Median |  | 3495 | 3494.645
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.01016%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 906b (Indicative Value 0.9 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.9 | 1
| 1 | 2 | 0.9 | 25
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 906b (Indicative Value 125 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 130 | 1
| 1 | 2 | 120 | 26
| Mean |  | 125 | 16
| Median |  | 125 | 125
| Std Dev. |  | 7.0711 | 32
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 906b (Indicative Value 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 27
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 906b (Indicative Value 18.3 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 29 | 1
| 1 | 2 | < 15 | 28
| Mean |  | 29 | 16
| Median |  | 29 | 18.2688
| Std Dev. |  |  | 34
Rel.Std.Dev.
| PDM3 |  | 58.74%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 906b (Indicative Value 0.32 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.319 | 29
| Mean |  | 0.3195 | 16
| Median |  | 0.3195 | 0.3195
| Std Dev. |  | 0.00070711 | 35
| Rel.Std.Dev. |  | 0.2213%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 906b (Indicative Value 7.21 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 7.22 | 1
| 1 | 2 | 7.206 | 22
| Mean |  | 7.213 | 16
| Median |  | 7.213 | 7.2128
| Std Dev. |  | 0.0098995 | 28
| Rel.Std.Dev. |  | 0.1372%
| PDM3 |  | 0.002267%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 906b (Indicative Value 3.53 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.52 | 1
| 1 | 2 | 3.53 | 23
| Mean |  | 3.525 | 16
| Median |  | 3.525 | 3.525
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.2006%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 906b (Indicative Value 0.405 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.41 | 1
| 1 | 2 | 0.4 | 24
| Mean |  | 0.405 | 16
| Median |  | 0.405 | 0.405
| Std Dev. |  | 0.0070711 | 30
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 906b (Indicative Value 0.04 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.04 | 25
| Mean |  | 0.04 | 16
| Median |  | 0.04 | 0.04
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 906b (Indicative Value 3.55 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.55 | 1
| 1 | 2 | 3.54 | 26
| Mean |  | 3.545 | 16
| Median |  | 3.545 | 3.545
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 0.1995%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 906b (Indicative Value 20 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 11
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 906b (Indicative Value 0.063 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0619 | 1
| 1 | 2 | 0.0642 | 28
| Mean |  | 0.06305 | 16
| Median |  | 0.06305 | 0.063019
| Std Dev. |  | 0.0016263 | 34
| Rel.Std.Dev. |  | 2.579%
| PDM3 |  | 0.04919%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 906b (Indicative Value 40 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 40 | 29
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 906b (Indicative Value 0.045 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.044 | 1
| 1 | 2 | 0.046 | 3
| Mean |  | 0.045 | 16
| Median |  | 0.045 | 0.045
| Std Dev. |  | 0.0014142 | 28
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 906b (Indicative Value 67 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 66.98 | 1
| 1 | 2 | 67.02 | 23
| Mean |  | 67 | 16
| Median |  | 67 | 67
| Std Dev. |  | 0.028284 | 29
| Rel.Std.Dev. |  | 0.04222%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 906b (Indicative Value < 10 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 10
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 906b (Indicative Value 170 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 170 | 25
| Mean |  | 170 | 16
| Median |  | 170 | 170
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 906b (Indicative Value 0.2 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 26
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 906b (Indicative Value 26.8 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 36 | 1
| 1 | 2 | 18 | 27
| Mean |  | 27 | 16
| Median |  | 27 | 26.778
| Std Dev. |  | 12.7279 | 33
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0.829%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 906b (Indicative Value 140 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 140 | 1
| 1 | 2 | 140 | 28
| Mean |  | 140 | 16
| Median |  | 140 | 140
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 906b (Indicative Value 285 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 290 | 1
| 1 | 2 | 280 | 14
| Mean |  | 285 | 16
| Median |  | 285 | 285
| Std Dev. |  | 7.0711 | 35
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 906b (Indicative Value 1.76 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.79 | 1
| 1 | 2 | 1.72 | 31
| Mean |  | 1.755 | 16
| Median |  | 1.755 | 1.755
| Std Dev. |  | 0.049497 | 37
| Rel.Std.Dev. |  | 2.82%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 906b (Indicative Value 0.7 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.7 | 9
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 906b (Indicative Value 12.6 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 12.6 | 10
| Mean |  | 12.6 | 16
| Median |  | 12.6 | 12.6
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 906b (Indicative Value 2765 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2780 | 1
| 1 | 2 | 2750 | 35
| Mean |  | 2765 | 16
| Median |  | 2765 | 2765
| Std Dev. |  | 21.2132 | 41
| Rel.Std.Dev. |  | 0.7672%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 906b (Indicative Value 3.3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.2 | 36
| Mean |  | 3.3 | 16
| Median |  | 3.3 | 3.3
| Std Dev. |  | 0.14142 | 42
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 906b (Indicative Value 10.3 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.2 | 1
| 1 | 2 | 10.3 | 37
| Mean |  | 10.25 | 16
| Median |  | 10.25 | 10.25
| Std Dev. |  | 0.070711 | 43
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 906b (Indicative Value 0.4 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 12
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 906b (Indicative Value 91 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 91.4 | 1
| 1 | 2 | 90.7 | 39
| Mean |  | 91.05 | 16
| Median |  | 91.05 | 91.05
| Std Dev. |  | 0.49497 | 45
| Rel.Std.Dev. |  | 0.5436%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 906b (Indicative Value 18.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.2 | 1
| 1 | 2 | 17.9 | 27
| Mean |  | 18.05 | 16
| Median |  | 18.05 | 18.05
| Std Dev. |  | 0.21213 | 46
| Rel.Std.Dev. |  | 1.175%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 906b (Indicative Value 11.5 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 13 | 13
| Mean |  | 11.5 | 16
| Median |  | 11.5 | 11.5
| Std Dev. |  | 2.1213 | 47
| Rel.Std.Dev. |  | 18.45%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 906b (Indicative Value 6.83 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.8 | 1
| 1 | 2 | 6.86 | 42
| Mean |  | 6.83 | 16
| Median |  | 6.83 | 6.83
| Std Dev. |  | 0.042426 | 48
| Rel.Std.Dev. |  | 0.6212%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 906b (Indicative Value 0.325 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.323 | 1
| 1 | 2 | 0.326 | 29
| Mean |  | 0.3245 | 16
| Median |  | 0.3245 | 0.3245
| Std Dev. |  | 0.0021213 | 49
| Rel.Std.Dev. |  | 0.6537%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 906b (Indicative Value 3.93 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.97 | 1
| 1 | 2 | 3.89 | 16
| Mean |  | 3.93 | 16
| Median |  | 3.93 | 3.93
| Std Dev. |  | 0.056569 | 50
| Rel.Std.Dev. |  | 1.439%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 906b (Indicative Value 1.13 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.1 | 1
| 1 | 2 | 1.15 | 17
| Mean |  | 1.125 | 16
| Median |  | 1.125 | 1.125
| Std Dev. |  | 0.035355 | 51
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 906b (Indicative Value 1.47 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.49 | 1
| 1 | 2 | 1.45 | 18
| Mean |  | 1.47 | 16
| Median |  | 1.47 | 1.47
| Std Dev. |  | 0.028284 | 52
| Rel.Std.Dev. |  | 1.924%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 906b (Indicative Value 27 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.8 | 1
| 1 | 2 | 27.2 | 47
| Mean |  | 27 | 16
| Median |  | 27 | 27
| Std Dev. |  | 0.28284 | 53
| Rel.Std.Dev. |  | 1.048%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 906b (Indicative Value 5.99 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.98 | 1
| 1 | 2 | 6 | 19
| Mean |  | 5.99 | 16
| Median |  | 5.99 | 5.99
| Std Dev. |  | 0.014142 | 54
| Rel.Std.Dev. |  | 0.2361%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 906b (Indicative Value 1.33 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.3 | 49
| Mean |  | 1.325 | 16
| Median |  | 1.325 | 1.325
| Std Dev. |  | 0.035355 | 55
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 906b (Indicative Value 7.71 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.67 | 1
| 1 | 2 | 7.74 | 33
| Mean |  | 7.705 | 16
| Median |  | 7.705 | 7.705
| Std Dev. |  | 0.049497 | 39
| Rel.Std.Dev. |  | 0.6424%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 906b (Indicative Value 0.56 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.56 | 1
| 1 | 2 | 0.55 | 16
| Mean |  | 0.555 | 16
| Median |  | 0.555 | 0.555
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.274%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 906b (Indicative Value 1.03 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1.05 | 35
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 906b (Indicative Value 45.9 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45.7 | 1
| 1 | 2 | 46 | 36
| Mean |  | 45.85 | 16
| Median |  | 45.85 | 45.85
| Std Dev. |  | 0.21213 | 42
| Rel.Std.Dev. |  | 0.4627%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 906b (Indicative Value 0.085 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.09 | 9
| Mean |  | 0.085 | 16
| Median |  | 0.085 | 0.085
| Std Dev. |  | 0.0070711 | 43
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 906b (Indicative Value 0.032 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0315 | 1
| 1 | 2 | 0.0321 | 38
| Mean |  | 0.0318 | 16
| Median |  | 0.0318 | 0.0318
| Std Dev. |  | 0.00042426 | 44
| Rel.Std.Dev. |  | 1.334%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 906b (Indicative Value 5 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 5.4 | 39
| Mean |  | 5 | 16
| Median |  | 5 | 5
| Std Dev. |  | 0.56569 | 45
| Rel.Std.Dev. |  | 11.31%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 906b (Indicative Value 19 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19 | 1
| 1 | 2 | 19 | 40
| Mean |  | 19 | 16
| Median |  | 19 | 19
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 906b (Indicative Value 38.5 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.7 | 1
| 1 | 2 | 38.3 | 18
| Mean |  | 38.5 | 16
| Median |  | 38.5 | 38.5
| Std Dev. |  | 0.28284 | 47
| Rel.Std.Dev. |  | 0.7347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 906b (Indicative Value 5 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 6 | 11
| Mean |  | 5 | 16
| Median |  | 5 | 5
| Std Dev. |  | 1.4142 | 48
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 906b (Indicative Value 31 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31 | 1
| 1 | 2 | 31 | 29
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 0 | 49
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 906b (Indicative Value 10.7 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.6 | 1
| 1 | 2 | 10.7 | 20
| Mean |  | 10.65 | 16
| Median |  | 10.65 | 10.65
| Std Dev. |  | 0.070711 | 50
| Rel.Std.Dev. |  | 0.664%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 906b (Indicative Value 133 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 132 | 1
| 1 | 2 | 133 | 45
| Mean |  | 132.5 | 16
| Median |  | 132.5 | 132.5
| Std Dev. |  | 0.70711 | 51
| Rel.Std.Dev. |  | 0.5337%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 906b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 12
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 52
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 906b (Indicative Value 2.1 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2 | 47
| Mean |  | 2.1 | 16
| Median |  | 2.1 | 2.1
| Std Dev. |  | 0.14142 | 53
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 906b (Indicative Value 3.95 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.9 | 1
| 1 | 2 | 4 | 14
| Mean |  | 3.95 | 16
| Median |  | 3.95 | 3.95
| Std Dev. |  | 0.070711 | 54
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 906b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 49
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 55
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 906b (Indicative Value 7.69 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.68 | 1
| 1 | 2 | 7.69 | 18
| Mean |  | 7.685 | 16
| Median |  | 7.685 | 7.685
| Std Dev. |  | 0.0070711 | 39
| Rel.Std.Dev. |  | 0.09201%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 906b (Indicative Value 8.1 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.2 | 1
| 1 | 2 | 8 | 10
| Mean |  | 8.1 | 16
| Median |  | 8.1 | 8.1
| Std Dev. |  | 0.14142 | 40
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 906b (Indicative Value 161 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 161 | 1
| 1 | 2 | 161 | 25
| Mean |  | 161 | 16
| Median |  | 161 | 161
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 906b (Indicative Value 1.48 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.45 | 1
| 1 | 2 | 1.51 | 36
| Mean |  | 1.48 | 16
| Median |  | 1.48 | 1.48
| Std Dev. |  | 0.042426 | 42
| Rel.Std.Dev. |  | 2.867%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 906b (Indicative Value 0.85 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.85 | 1
| 1 | 2 | 0.84 | 37
| Mean |  | 0.845 | 16
| Median |  | 0.845 | 0.845
| Std Dev. |  | 0.0070711 | 43
| Rel.Std.Dev. |  | 0.8368%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 906b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 11
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 906b (Indicative Value 14.6 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.5 | 1
| 1 | 2 | 14.6 | 39
| Mean |  | 14.55 | 16
| Median |  | 14.55 | 14.55
| Std Dev. |  | 0.070711 | 45
| Rel.Std.Dev. |  | 0.486%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 906b (Indicative Value 0.119 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.118 | 1
| 1 | 2 | 0.119 | 40
| Mean |  | 0.1185 | 16
| Median |  | 0.1185 | 0.1185
| Std Dev. |  | 0.00070711 | 46
| Rel.Std.Dev. |  | 0.5967%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 906b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 41
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 906b (Indicative Value 0.14 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.13 | 1
| 1 | 2 | 0.14 | 19
| Mean |  | 0.135 | 16
| Median |  | 0.135 | 0.135
| Std Dev. |  | 0.0070711 | 48
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 906b (Indicative Value 5.13 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.13 | 1
| 1 | 2 | 5.13 | 43
| Mean |  | 5.13 | 16
| Median |  | 5.13 | 5.13
| Std Dev. |  | 0 | 49
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 906b (Indicative Value 4.8 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 4.8 | 12
| Mean |  | 4.8 | 16
| Median |  | 4.8 | 4.8
| Std Dev. |  | 0 | 50
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 906b (Indicative Value 3.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 3.5 | 13
| Mean |  | 3.75 | 16
| Median |  | 3.75 | 3.75
| Std Dev. |  | 0.35355 | 51
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 906b (Indicative Value 16.1 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.1 | 1
| 1 | 2 | 16.1 | 46
| Mean |  | 16.1 | 16
| Median |  | 16.1 | 16.1
| Std Dev. |  | 0 | 52
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 906b (Indicative Value 0.71 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.71 | 1
| 1 | 2 | 0.7 | 47
| Mean |  | 0.705 | 16
| Median |  | 0.705 | 0.705
| Std Dev. |  | 0.0070711 | 53
| Rel.Std.Dev. |  | 1.003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 906b (Indicative Value 153 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 150 | 1
| 1 | 2 | 155 | 28
| Mean |  | 152.5 | 16
| Median |  | 152.5 | 152.5
| Std Dev. |  | 3.5355 | 54
| Rel.Std.Dev. |  | 2.318%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 906b (Indicative Value 284 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 282 | 1
| 1 | 2 | 285 | 14
| Mean |  | 283.5 | 16
| Median |  | 283.5 | 283.5
| Std Dev. |  | 2.1213 | 55
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.