## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 550b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 0.32344 | 0.24826 | 0.39861 | 0.25679 | 0.39009
Al, Aluminium (wt.%) | 5.8761 | 5.7317 | 6.0205 | 5.77 | 5.9822
As, Arsenic (ppm) | 60.5426 | 57.9303 | 63.155 | 58.6541 | 62.4311
Ba, Barium (ppm) | 735.3056 | 709.0075 | 761.6036 | 710.2913 | 760.3199
Be, Beryllium (ppm) | 3.3393 | 3.1743 | 3.5042 | 3.2529 | 3.4257
Bi, Bismuth (ppm) | 2.1024 | 1.9953 | 2.2096 | 2.0358 | 2.169
Ca, Calcium (wt.%) | 1.3386 | 1.3063 | 1.3709 | 1.3057 | 1.3715
Cd, Cadmium (ppm) | 0.20152 | 0.17623 | 0.22681 | 0.17365 | 0.22939
Ce, Cerium (ppm) | 120.0018 | 113.4631 | 126.5405 | 117.1404 | 122.8632
Co, Cobalt (wt.%) | 0.14306 | 0.13891 | 0.14721 | 0.14086 | 0.14525
Cr, Chromium (ppm) | 71.9608 | 67.9021 | 76.0195 | 68.5706 | 75.351
Cs, Caesium (ppm) | 2.2165 | 2.1187 | 2.3143 | 2.1349 | 2.2981
Cu, Copper (wt.%) | 1.8401 | 1.8108 | 1.8694 | 1.8215 | 1.8587
Dy, Dysprosium (ppm) | 3.8059 | 3.4118 | 4.1999 | 3.6852 | 3.9265
Er, Erbium (ppm) | 2.1789 | 1.9644 | 2.3933 | 2.0922 | 2.2655
Eu, Europium (ppm) | 0.82332 | 0.75304 | 0.89359 | 0.76576 | 0.88087
Fe, Iron (wt.%) | 2.7398 | 2.6551 | 2.8245 | 2.6863 | 2.7933
Ga, Gallium (ppm) | 22.7422 | 21.8995 | 23.5848 | 22.1728 | 23.3115
Gd, Gadolinium (ppm) | 4.6715 | 4.1103 | 5.2327 | 4.5043 | 4.8387
Hf, Hafnium (ppm) | 3.8276 | 3.6597 | 3.9955 | 3.6537 | 4.0015
Ho, Holmium (ppm) | 0.73479 | 0.60291 | 0.86668 | 0.70301 | 0.76657
In, Indium (ppm) | 0.96135 | 0.91943 | 1.0033 | 0.93248 | 0.99023
K, Potassium (wt.%) | 2.0174 | 1.9699 | 2.0648 | 1.969 | 2.0658
La, Lanthanum (ppm) | 60.6183 | 56.5846 | 64.652 | 59.1495 | 62.0871
Li, Lithium (ppm) | 76.9276 | 74.8397 | 79.0156 | 75.0907 | 78.7645
Lu, Lutetium (ppm) | 0.33275 | 0.30307 | 0.36243 | 0.30851 | 0.35699
Mg, Magnesium (wt.%) | 4.3121 | 4.194 | 4.4303 | 4.217 | 4.4073
Mn, Manganese (wt.%) | 0.019058 | 0.01849 | 0.019626 | 0.018417 | 0.019699
Mo, Molybdenum (ppm) | 6.9731 | 6.7132 | 7.2331 | 6.7664 | 7.1799
Na, Sodium (wt.%) | 0.052798 | 0.049203 | 0.056393 | 0.050674 | 0.054922
Nb, Niobium (ppm) | 10.7128 | 9.1757 | 12.2499 | 10.2827 | 11.1429
Nd, Neodymium (ppm) | 48.4473 | 44.7665 | 52.128 | 47.0091 | 49.8855
Ni, Nickel (ppm) | 67.2048 | 64.9948 | 69.4148 | 65.5222 | 68.8874
P, Phosphorus (wt.%) | 0.049977 | 0.047718 | 0.052236 | 0.048462 | 0.051492
Pb, Lead (ppm) | 13.1704 | 12.3549 | 13.9859 | 12.51 | 13.8308
Pr, Praseodymium (ppm) | 13.7721 | 12.8886 | 14.6556 | 13.3133 | 14.2309
Rb, Rubidium (ppm) | 88.2447 | 84.9892 | 91.5002 | 85.5263 | 90.9631
Re, Rhenium (ppm) | 0.013526 | 0.011179 | 0.015874 | IND | IND
S, Sulphur (wt.%) | 1.1574 | 1.1229 | 1.1918 | 1.1354 | 1.1793
Sb, Antimony (ppm) | 2.6159 | 2.4039 | 2.8279 | 2.4674 | 2.7644
Sc, Scandium (ppm) | 11.3035 | 10.8015 | 11.8055 | 11.0393 | 11.5676
Se, Selenium (ppm) | 3.6942 | 2.8387 | 4.5497 | 3.3121 | 4.0763
Sm, Samarium (ppm) | 7.3981 | 6.6769 | 8.1193 | 7.1182 | 7.678
Sn, Tin (ppm) | 2.4713 | 2.325 | 2.6177 | 2.3362 | 2.6065
Sr, Strontium (ppm) | 77.4751 | 74.9591 | 79.991 | 75.7575 | 79.1926
Ta, Tantalum (ppm) | 0.74095 | 0.62546 | 0.85644 | 0.69218 | 0.78971
Tb, Terbium (ppm) | 0.6795 | 0.61263 | 0.74637 | 0.6507 | 0.7083
Te, Tellurium (ppm) | 0.07025 | 0.047456 | 0.093044 | IND | IND
Th, Thorium (ppm) | 13.3329 | 12.6429 | 14.023 | 12.8943 | 13.7715
Ti, Titanium (wt.%) | 0.21423 | 0.18768 | 0.24078 | 0.20722 | 0.22124
Tl, Thallium (ppm) | 0.53653 | 0.50725 | 0.56581 | 0.51524 | 0.55782
Tm, Thulium (ppm) | 0.33177 | 0.28873 | 0.37482 | 0.31271 | 0.35083
U, Uranium (ppm) | 8.8288 | 8.3953 | 9.2622 | 8.5993 | 9.0583
V, Vanadium (ppm) | 200.6645 | 193.8025 | 207.5265 | 195.5435 | 205.7855
W, Tungsten (ppm) | 2.7063 | 2.3274 | 3.0852 | 2.4311 | 2.9814
Y, Yttrium (ppm) | 19.0781 | 18.2255 | 19.9307 | 18.6305 | 19.5257
Yb, Ytterbium (ppm) | 2.2163 | 2.0759 | 2.3567 | 2.1363 | 2.2963
Zn, Zinc (ppm) | 26.6105 | 25.0329 | 28.188 | 25.2645 | 27.9564
Zr, Zirconium (ppm) | 134.4601 | 129.7168 | 139.2033 | 131.4204 | 137.4998
Sulphuric Acid 5% Leach
Co, Cobalt (wt.%) | 0.042001 | 0.038344 | 0.045658 | 0.040816 | 0.043186
Cu, Copper (wt.%) | 0.83217 | 0.80543 | 0.85891 | 0.81295 | 0.85138
Infrared Combustion
C, Carbon (wt.%) | 3.1219 | 3.0465 | 3.1973 | 3.0917 | 3.152
S, Sulphur (wt.%) | 1.1657 | 1.1362 | 1.1951 | 1.1431 | 1.1882
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 11.4037 | 11.2181 | 11.5892 | 11.2893 | 11.518
BaO, Barium oxide (ppm) | 824.1849 | 736.5115 | 911.8582 | IND | IND
CaO, Calcium oxide (wt.%) | 1.8724 | 1.8462 | 1.8985 | 1.8494 | 1.8954
Co, Cobalt (wt.%) | 0.14455 | 0.13932 | 0.14978 | 0.14209 | 0.14701
Cu, Copper (wt.%) | 1.8423 | 1.8061 | 1.8785 | 1.8136 | 1.871
Fe2O3, Iron(III) oxide (wt.%) | 4.0116 | 3.9492 | 4.0739 | 3.9506 | 4.0726
K2O, Potassium oxide (wt.%) | 2.4195 | 2.3908 | 2.4483 | 2.3894 | 2.4497
MgO, Magnesium oxide (wt.%) | 7.1901 | 7.0525 | 7.3276 | 7.1343 | 7.2458
MnO, Manganese oxide (wt.%) | 0.025504 | 0.023202 | 0.027805 | 0.023616 | 0.027391
Ni, Nickel (ppm) | 56.6781 | 28.2783 | 85.0779 | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.11534 | 0.11034 | 0.12034 | 0.11241 | 0.11827
SiO2, Silicon dioxide (wt.%) | 58.1847 | 57.5711 | 58.7982 | 57.7759 | 58.5934
SO3, Sulphur trioxide (wt.%) | 3.0023 | 2.879 | 3.1256 | 2.9232 | 3.0814
SrO, Strontium oxide (ppm) | 102.1716 | 87.0328 | 117.3104 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.64592 | 0.62839 | 0.66345 | 0.63195 | 0.65989
V2O5, Vanadium(V) oxide (ppm) | 362.5762 | 285.0772 | 440.0752 | IND | IND
Zr, Zirconium (ppm) | 184.2244 | 140.1716 | 228.2773 | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 9.9748 | 9.6689 | 10.2807 | 9.8562 | 10.0933
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.8847 | 5.7563 | 6.0131 | 5.7838 | 5.9856
As, Arsenic (ppm) | 59.9998 | 53.7826 | 66.2171 | 57.9539 | 62.0457
B, Boron (ppm) | 170.7019 | 147.099 | 194.3048 | 162.0546 | 179.3491
Ba, Barium (ppm) | 735.4886 | 705.5412 | 765.436 | 719.6462 | 751.331
Be, Beryllium (ppm) | 3.76 | 2.8811 | 4.6389 | IND | IND
Bi, Bismuth (ppm) | 2.1547 | 1.9095 | 2.3999 | 2.0158 | 2.2935
Ca, Calcium (wt.%) | 1.3528 | 1.2969 | 1.4086 | 1.3179 | 1.3876
Ce, Cerium (ppm) | 126.4803 | 120.4335 | 132.5271 | 122.547 | 130.4136
Co, Cobalt (wt.%) | 0.14509 | 0.13997 | 0.1502 | 0.14234 | 0.14783
Cr, Chromium (ppm) | 87.1448 | 73.7421 | 100.5476 | 82.6429 | 91.6468
Cs, Caesium (ppm) | 2.2573 | 2.0854 | 2.4291 | 2.0933 | 2.4212
Cu, Copper (wt.%) | 1.8305 | 1.7867 | 1.8743 | 1.7978 | 1.8631
Dy, Dysprosium (ppm) | 5.7834 | 5.3446 | 6.2223 | 5.5627 | 6.0042
Er, Erbium (ppm) | 3.1387 | 2.7774 | 3.4999 | 2.8907 | 3.3867
Eu, Europium (ppm) | 0.92213 | 0.7929 | 1.0514 | 0.83397 | 1.0103
Fe, Iron (wt.%) | 2.7424 | 2.6704 | 2.8143 | 2.694 | 2.7907
Ga, Gallium (ppm) | 23.0023 | 21.6299 | 24.3747 | 21.7034 | 24.3013
Gd, Gadolinium (ppm) | 6.1183 | 5.6101 | 6.6266 | 5.7384 | 6.4983
Ge, Germanium (ppm) | 2.01 | 1.696 | 2.324 | IND | IND
Ho, Holmium (ppm) | 1.1428 | 1.0264 | 1.2591 | 1.0638 | 1.2217
In, Indium (ppm) | 0.98272 | 0.83386 | 1.1316 | IND | IND
K, Potassium (wt.%) | 2.0317 | 1.9607 | 2.1026 | 1.9647 | 2.0986
La, Lanthanum (ppm) | 65.9656 | 63.3337 | 68.5976 | 63.9801 | 67.9512
Li, Lithium (ppm) | 78.3646 | 75.6172 | 81.1121 | 76.0954 | 80.6338
Lu, Lutetium (ppm) | 0.47695 | 0.41219 | 0.54171 | 0.438 | 0.5159
Mg, Magnesium (wt.%) | 4.3484 | 4.2356 | 4.4611 | 4.2656 | 4.4312
Mn, Manganese (wt.%) | 0.019522 | 0.018647 | 0.020397 | 0.019185 | 0.019859
Mo, Molybdenum (ppm) | 7.0692 | 6.2194 | 7.9191 | IND | IND
Nb, Niobium (ppm) | 19.4534 | 18.1605 | 20.7463 | 18.0855 | 20.8213
Nd, Neodymium (ppm) | 52.6152 | 50.6131 | 54.6173 | 50.4213 | 54.8091
Ni, Nickel (ppm) | 69.5639 | 59.4154 | 79.7124 | 65.9387 | 73.1892
P, Phosphorus (wt.%) | 0.05178 | 0.045797 | 0.057764 | IND | IND
Pb, Lead (ppm) | 14.9633 | 11.933 | 17.9937 | IND | IND
Pr, Praseodymium (ppm) | 14.4294 | 13.3381 | 15.5207 | 13.7456 | 15.1132
Rb, Rubidium (ppm) | 87.7846 | 81.3611 | 94.2082 | 85.1919 | 90.3773
Re, Rhenium (ppm) | < 0.1 | IND | IND | IND | IND
S, Sulphur (wt.%) | 1.179 | 1.1244 | 1.2337 | 1.1507 | 1.2073
Sb, Antimony (ppm) | 2.6653 | 2.2723 | 3.0583 | 2.2694 | 3.0612
Sc, Scandium (ppm) | 11.12 | 10.2536 | 11.9864 | IND | IND
Si, Silicon (wt.%) | 27.2817 | 26.6075 | 27.956 | 26.8521 | 27.7114
Sm, Samarium (ppm) | 7.6994 | 7.1979 | 8.201 | 7.334 | 8.0649
Sn, Tin (ppm) | 3.0195 | 2.2234 | 3.8155 | IND | IND
Sr, Strontium (ppm) | 78.2221 | 73.5588 | 82.8855 | 75.6789 | 80.7653
Ta, Tantalum (ppm) | 1.5326 | 1.2984 | 1.7668 | 1.2042 | 1.8611
Tb, Terbium (ppm) | 0.96125 | 0.89124 | 1.0313 | 0.9065 | 1.016
Th, Thorium (ppm) | 13.6757 | 13.1443 | 14.207 | 13.2526 | 14.0988
Ti, Titanium (wt.%) | 0.37423 | 0.36094 | 0.38752 | 0.36346 | 0.385
Tl, Thallium (ppm) | 0.59784 | 0.50503 | 0.69064 | IND | IND
Tm, Thulium (ppm) | 0.48575 | 0.42807 | 0.54343 | 0.45739 | 0.51411
U, Uranium (ppm) | 9.117 | 8.7101 | 9.5239 | 8.85 | 9.384
V, Vanadium (ppm) | 213.6536 | 204.5768 | 222.7304 | 208.835 | 218.4723
W, Tungsten (ppm) | 3.4021 | 2.7306 | 4.0736 | IND | IND
Y, Yttrium (ppm) | 29.4088 | 27.8839 | 30.9337 | 27.8352 | 30.9824
Yb, Ytterbium (ppm) | 3.1314 | 2.8645 | 3.3984 | 2.8525 | 3.4103
Aqua Regia Digestion
Ag, Silver (ppm) | 0.23312 | 0.17639 | 0.28985 | 0.20686 | 0.25938
Al, Aluminium (wt.%) | 1.3109 | 1.2475 | 1.3743 | 1.2677 | 1.3541
As, Arsenic (ppm) | 58.4178 | 56.3532 | 60.4824 | 56.8424 | 59.9932
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 103.4893 | 96.1253 | 110.8534 | 98.9278 | 108.0509
Be, Beryllium (ppm) | 1.6285 | 1.5199 | 1.7371 | 1.5857 | 1.6713
Bi, Bismuth (ppm) | 1.9093 | 1.826 | 1.9927 | 1.8556 | 1.9631
Ca, Calcium (wt.%) | 1.3184 | 1.2834 | 1.3533 | 1.2938 | 1.3429
Cd, Cadmium (ppm) | 0.19621 | 0.18006 | 0.21236 | 0.1763 | 0.21612
Ce, Cerium (ppm) | 43.3083 | 40.3665 | 46.2502 | 41.4235 | 45.1931
Co, Cobalt (wt.%) | 0.13079 | 0.12498 | 0.1366 | 0.12809 | 0.13349
Cr, Chromium (ppm) | 28.4997 | 27.1082 | 29.8911 | 27.157 | 29.8424
Cs, Caesium (ppm) | 0.4544 | 0.4105 | 0.4983 | 0.43673 | 0.47207
Cu, Copper (wt.%) | 1.8171 | 1.7824 | 1.8518 | 1.7966 | 1.8376
Dy, Dysprosium (ppm) | 1.1679 | 1.0202 | 1.3156 | 1.126 | 1.2097
Er, Erbium (ppm) | 0.51885 | 0.47499 | 0.5627 | 0.49271 | 0.54498
Eu, Europium (ppm) | 0.39055 | 0.3343 | 0.44681 | 0.37324 | 0.40786
Fe, Iron (wt.%) | 2.4576 | 2.3945 | 2.5207 | 2.4133 | 2.502
Ga, Gallium (ppm) | 7.3108 | 6.7985 | 7.8232 | 7.044 | 7.5777
Gd, Gadolinium (ppm) | 2.1919 | 1.9374 | 2.4464 | 2.1035 | 2.2803
Hf, Hafnium (ppm) | 0.32104 | 0.29487 | 0.34721 | 0.29966 | 0.34242
Hg, Mercury (ppm) | 0.047214 | 0.034249 | 0.06018 | IND | IND
Ho, Holmium (ppm) | 0.19551 | 0.17255 | 0.21846 | 0.18159 | 0.20942
In, Indium (ppm) | 0.62623 | 0.59736 | 0.6551 | 0.60789 | 0.64457
K, Potassium (wt.%) | 0.20498 | 0.19346 | 0.21651 | 0.1967 | 0.21327
La, Lanthanum (ppm) | 19.2317 | 18.1145 | 20.3489 | 18.5995 | 19.8638
Li, Lithium (ppm) | 49.8796 | 45.9894 | 53.7699 | 48.0499 | 51.7093
Lu, Lutetium (ppm) | 0.065362 | 0.059319 | 0.071405 | 0.060552 | 0.070171
Mg, Magnesium (wt.%) | 3.5321 | 3.3957 | 3.6686 | 3.4552 | 3.6091
Mn, Manganese (wt.%) | 0.017955 | 0.017327 | 0.018582 | 0.017486 | 0.018423
Mo, Molybdenum (ppm) | 6.5632 | 6.3327 | 6.7937 | 6.373 | 6.7534
Na, Sodium (wt.%) | 0.0095787 | 0.0085495 | 0.010608 | IND | IND
Nd, Neodymium (ppm) | 17.4893 | 13.8158 | 21.1628 | 17.2566 | 17.722
Ni, Nickel (ppm) | 60.0636 | 57.6273 | 62.4999 | 58.2647 | 61.8624
P, Phosphorus (wt.%) | 0.039341 | 0.037782 | 0.040901 | 0.038048 | 0.040635
Pb, Lead (ppm) | 9.1344 | 8.4427 | 9.826 | 8.753 | 9.5158
Pr, Praseodymium (ppm) | 4.6782 | 3.7478 | 5.6086 | 4.5919 | 4.7646
Rb, Rubidium (ppm) | 8.7004 | 8.0285 | 9.3724 | 8.3675 | 9.0333
Re, Rhenium (ppm) | 0.011667 | 0.0098187 | 0.013515 | IND | IND
S, Sulphur (wt.%) | 1.1612 | 1.1276 | 1.1948 | 1.1427 | 1.1796
Sb, Antimony (ppm) | 1.6192 | 1.4645 | 1.7738 | 1.5207 | 1.7177
Sc, Scandium (ppm) | 3.6846 | 3.5066 | 3.8627 | 3.5278 | 3.8415
Se, Selenium (ppm) | 3.5038 | 3.1298 | 3.8778 | 3.2124 | 3.7952
Sm, Samarium (ppm) | 3.0035 | 2.4952 | 3.5119 | 2.8976 | 3.1094
Sn, Tin (ppm) | 0.83287 | 0.78569 | 0.88006 | 0.79491 | 0.87083
Sr, Strontium (ppm) | 29.6478 | 28.2475 | 31.0481 | 28.6745 | 30.6211
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.25977 | 0.23999 | 0.27955 | 0.24729 | 0.27225
Te, Tellurium (ppm) | 0.048338 | 0.034956 | 0.06172 | IND | IND
Th, Thorium (ppm) | 6.1023 | 5.6935 | 6.5111 | 5.8636 | 6.341
U, Uranium (ppm) | 4.6116 | 4.3864 | 4.8368 | 4.4915 | 4.7317
V, Vanadium (ppm) | 45.9389 | 42.8889 | 48.9888 | 44.5056 | 47.3722
W, Tungsten (ppm) | 1.0022 | 0.76688 | 1.2374 | 0.87094 | 1.1334
Y, Yttrium (ppm) | 4.5461 | 4.3199 | 4.7723 | 4.3993 | 4.6929
Yb, Ytterbium (ppm) | 0.42848 | 0.39341 | 0.46354 | 0.40037 | 0.45658
Zn, Zinc (ppm) | 22.872 | 21.7142 | 24.0299 | 21.742 | 24.002
Zr, Zirconium (ppm) | 10.9703 | 10.2676 | 11.673 | 10.5753 | 11.3652
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 550b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 18.5 | Hg | ppm | 0.05625
Ge | ppm | 0.23053 | Pt | ppb | 7.8333
Borate Fusion XRF
Ag | ppm | 0.0385 | Hf | ppm | < 80 | Se | ppm | 3.5
As | ppm | 85.6667 | Ho | ppm | 1.08 | Sm | ppm | 7.05
Be | ppm | 2 | In | ppm | 0.98333 | Sn | ppm | < 50
Bi | ppm | < 100 | La | ppm | 72.8341 | Ta | ppm | 34.7972
Cd | ppm | < 10 | Lu | ppm | 0.41833 | Tb | ppm | 0.90167
Ce | ppm | 122.5779 | Mo | ppm | 16.7667 | Te | ppm | < 0.1
Cl | ppm | 350 | Na2O | wt.% | 0.083571 | Th | ppm | 13.8117
Cr2O3 | ppm | 117.533 | Nb | ppm | 25.39 | Tl | ppm | 0.465
Cs | ppm | 2.1167 | Nd | ppm | 49.445 | Tm | ppm | 0.44333
Dy | ppm | 5.2817 | Pb | ppm | 48.2458 | U | ppm | 7.2033
Er | ppm | 2.99 | Pr | ppm | 14.8308 | W | ppm | < 10
Eu | ppm | 0.85667 | Rb | ppm | 59.5333 | Y | ppm | 38.8834
F | ppm | < 5000 | Re | ppm | < 0.1 | Yb | ppm | 2.85
Ga | ppm | 22.3 | Sb | ppm | < 50 | Zn | ppm | < 50
Gd | ppm | 5.825 | Sc | ppm | 11.6667
Peroxide Fusion ICP
Ag | ppm | 0.6325 | Hg | ppm | < 5 | Te | ppm | < 1
Cd | ppm | < 10 | Na | wt.% | 0.039216 | Zn | ppm | 31.4668
Hf | ppm | 6.291 | Se | ppm | 7.5833 | Zr | ppm | 174.1111
Aqua Regia Digestion
Ge | ppm | 0.077083 | Pt | ppb | 1.6667 | Tm | ppm | 0.071142
Nb | ppm | 0.048032 | Ti | wt.% | 0.0034021
Pd | ppb | < 10 | Tl | ppm | 0.17627
3-Acid Digestion (no HF)
As | ppm | < 10 | Ni | ppm | 72.5 | V | ppm | 71.6667
Co | wt.% | 0.13517 | Pb | ppm | 52.5 | Zn | ppm | 75
Cu | wt.% | 1.8133 | S | wt.% | 1.1567
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 550b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 0.32344 | 0.052649 | 0.21814 | 0.42873 | 0.16549 | 0.48138 | 16.28% | 32.56% | 48.83% | 0.30726 | 0.33961
Al, wt.% | 5.8761 | 0.19007 | 5.4959 | 6.2562 | 5.3059 | 6.4463 | 3.235% | 6.469% | 9.704% | 5.5823 | 6.1699
As, ppm | 60.5426 | 3.1736 | 54.1954 | 66.8898 | 51.0218 | 70.0634 | 5.242% | 10.48% | 15.73% | 57.5155 | 63.5697
Ba, ppm | 735.3056 | 41.5202 | 652.2651 | 818.346 | 610.7449 | 859.8662 | 5.647% | 11.29% | 16.94% | 698.5403 | 772.0709
Be, ppm | 3.3393 | 0.21321 | 2.9128 | 3.7657 | 2.6996 | 3.9789 | 6.385% | 12.77% | 19.16% | 3.1723 | 3.5062
Bi, ppm | 2.1024 | 0.1195 | 1.8634 | 2.3414 | 1.7439 | 2.4609 | 5.684% | 11.37% | 17.05% | 1.9973 | 2.2076
Ca, wt.% | 1.3386 | 0.031846 | 1.2749 | 1.4023 | 1.243 | 1.4341 | 2.379% | 4.758% | 7.137% | 1.2716 | 1.4055
Cd, ppm | 0.20152 | 0.014662 | 0.17219 | 0.23084 | 0.15753 | 0.2455 | 7.276% | 14.55% | 21.83% | 0.19144 | 0.21159
Ce, ppm | 120.0018 | 8.7933 | 102.4152 | 137.5884 | 93.6219 | 146.3817 | 7.328% | 14.66% | 21.98% | 114.0017 | 126.0019
Co, wt.% | 0.14306 | 0.0048829 | 0.13329 | 0.15282 | 0.12841 | 0.15771 | 3.413% | 6.827% | 10.24% | 0.1359 | 0.15021
Cr, ppm | 71.9608 | 6.2798 | 59.4012 | 84.5204 | 53.1214 | 90.8002 | 8.727% | 17.45% | 26.18% | 68.3628 | 75.5589
Cs, ppm | 2.2165 | 0.10758 | 2.0014 | 2.4317 | 1.8938 | 2.5393 | 4.854% | 9.707% | 14.56% | 2.1057 | 2.3273
Cu, wt.% | 1.8401 | 0.032273 | 1.7756 | 1.9047 | 1.7433 | 1.9369 | 1.754% | 3.508% | 5.262% | 1.7481 | 1.9321
Dy, ppm | 3.8059 | 0.28756 | 3.2307 | 4.381 | 2.9432 | 4.6685 | 7.556% | 15.11% | 22.67% | 3.6156 | 3.9962
Er, ppm | 2.1789 | 0.18397 | 1.8109 | 2.5468 | 1.6269 | 2.7308 | 8.443% | 16.89% | 25.33% | 2.0699 | 2.2878
Eu, ppm | 0.82332 | 0.053321 | 0.71667 | 0.92996 | 0.66335 | 0.98328 | 6.476% | 12.95% | 19.43% | 0.78215 | 0.86448
Fe, wt.% | 2.7398 | 0.082204 | 2.5754 | 2.9042 | 2.4932 | 2.9864 | 3% | 6.001% | 9.001% | 2.6028 | 2.8768
Ga, ppm | 22.7422 | 1.1355 | 20.4711 | 25.0132 | 19.3356 | 26.1487 | 4.993% | 9.986% | 14.98% | 21.6051 | 23.8793
Gd, ppm | 4.6715 | 0.42915 | 3.8132 | 5.5298 | 3.384 | 5.9589 | 9.187% | 18.37% | 27.56% | 4.4379 | 4.9051
Hf, ppm | 3.8276 | 0.18506 | 3.4575 | 4.1977 | 3.2724 | 4.3828 | 4.835% | 9.67% | 14.5% | 3.6362 | 4.019
Ho, ppm | 0.73479 | 0.13118 | 0.47242 | 0.99716 | 0.34124 | 1.1283 | 17.85% | 35.71% | 53.56% | 0.69805 | 0.77153
In, ppm | 0.96135 | 0.048652 | 0.86405 | 1.0587 | 0.8154 | 1.1073 | 5.061% | 10.12% | 15.18% | 0.91329 | 1.0094
K, wt.% | 2.0174 | 0.046789 | 1.9238 | 2.111 | 1.877 | 2.1578 | 2.319% | 4.639% | 6.958% | 1.9165 | 2.1183
La, ppm | 60.6183 | 5.7172 | 49.1838 | 72.0527 | 43.4666 | 77.77 | 9.432% | 18.86% | 28.29% | 57.5874 | 63.6492
Li, ppm | 76.9276 | 2.7761 | 71.3754 | 82.4798 | 68.5993 | 85.2559 | 3.609% | 7.217% | 10.83% | 73.0812 | 80.774
Lu, ppm | 0.33275 | 0.01738 | 0.29799 | 0.36751 | 0.28061 | 0.38489 | 5.223% | 10.45% | 15.67% | 0.31611 | 0.34939
Mg, wt.% | 4.3121 | 0.12865 | 4.0548 | 4.5695 | 3.9262 | 4.6981 | 2.984% | 5.967% | 8.951% | 4.0965 | 4.5278
Mn, wt.% | 0.019058 | 0.00060172 | 0.017855 | 0.020261 | 0.017253 | 0.020863 | 3.157% | 6.315% | 9.472% | 0.018105 | 0.020011
Mo, ppm | 6.9731 | 0.34609 | 6.2809 | 7.6653 | 5.9349 | 8.0114 | 4.963% | 9.926% | 14.89% | 6.6245 | 7.3218
Na, wt.% | 0.052798 | 0.0070176 | 0.038763 | 0.066833 | 0.031745 | 0.073851 | 13.29% | 26.58% | 39.87% | 0.050158 | 0.055438
Nb, ppm | 10.7128 | 3.1784 | 4.3559 | 17.0696 | 1.1775 | 20.2481 | 29.67% | 59.34% | 89.01% | 10.1771 | 11.2484
Nd, ppm | 48.4473 | 2.9507 | 42.5458 | 54.3487 | 39.5951 | 57.2995 | 6.091% | 12.18% | 18.27% | 46.0249 | 50.8696
Ni, ppm | 67.2048 | 1.9347 | 63.3355 | 71.0741 | 61.4008 | 73.0088 | 2.879% | 5.758% | 8.636% | 63.8446 | 70.565
P, wt.% | 0.049977 | 0.0022865 | 0.045404 | 0.05455 | 0.043117 | 0.056836 | 4.575% | 9.15% | 13.73% | 0.047478 | 0.052475
Pb, ppm | 13.1704 | 1.3316 | 10.5072 | 15.8337 | 9.1755 | 17.1653 | 10.11% | 20.22% | 30.33% | 12.5119 | 13.8289
Pr, ppm | 13.7721 | 0.75908 | 12.2539 | 15.2902 | 11.4949 | 16.0493 | 5.512% | 11.02% | 16.54% | 13.0835 | 14.4607
Rb, ppm | 88.2447 | 4.3009 | 79.6428 | 96.8466 | 75.3419 | 101.1475 | 4.874% | 9.748% | 14.62% | 83.8325 | 92.6569
Re, ppm | 0.013526 | 0.001535 | 0.010456 | 0.016596 | 0.0089213 | 0.018131 | 11.35% | 22.7% | 34.04% | 0.01285 | 0.014203
S, wt.% | 1.1574 | 0.041022 | 1.0753 | 1.2394 | 1.0343 | 1.2804 | 3.544% | 7.089% | 10.63% | 1.0995 | 1.2152
Sb, ppm | 2.6159 | 0.22847 | 2.159 | 3.0729 | 1.9305 | 3.3013 | 8.734% | 17.47% | 26.2% | 2.4851 | 2.7467
Sc, ppm | 11.3035 | 0.57822 | 10.147 | 12.4599 | 9.5688 | 13.0381 | 5.115% | 10.23% | 15.35% | 10.7383 | 11.8686
Se, ppm | 3.6942 | 0.52131 | 2.6516 | 4.7368 | 2.1303 | 5.2581 | 14.11% | 28.22% | 42.33% | 3.5095 | 3.8789
Sm, ppm | 7.3981 | 0.70431 | 5.9895 | 8.8067 | 5.2852 | 9.5111 | 9.52% | 19.04% | 28.56% | 7.0282 | 7.768
Sn, ppm | 2.4713 | 0.17459 | 2.1221 | 2.8205 | 1.9476 | 2.9951 | 7.065% | 14.13% | 21.19% | 2.3478 | 2.5949
Sr, ppm | 77.4751 | 3.2161 | 71.0429 | 83.9072 | 67.8268 | 87.1233 | 4.151% | 8.302% | 12.45% | 73.6013 | 81.3488
Ta, ppm | 0.74095 | 0.20134 | 0.33827 | 1.1436 | 0.13693 | 1.345 | 27.17% | 54.35% | 81.52% | 0.7039 | 0.77799
Tb, ppm | 0.6795 | 0.040139 | 0.59922 | 0.75978 | 0.55909 | 0.79992 | 5.907% | 11.81% | 17.72% | 0.64553 | 0.71348
Te, ppm | 0.07025 | 0.017186 | 0.035879 | 0.10462 | 0.018693 | 0.12181 | 24.46% | 48.93% | 73.39% | 0.066737 | 0.073762
Th, ppm | 13.3329 | 0.68091 | 11.9711 | 14.6947 | 11.2902 | 15.3757 | 5.107% | 10.21% | 15.32% | 12.6663 | 13.9996
Ti, wt.% | 0.21423 | 0.052922 | 0.10839 | 0.32007 | 0.055466 | 0.373 | 24.7% | 49.41% | 74.11% | 0.20352 | 0.22494
Tl, ppm | 0.53653 | 0.046127 | 0.44427 | 0.62878 | 0.39815 | 0.67491 | 8.597% | 17.19% | 25.79% | 0.5097 | 0.56336
Tm, ppm | 0.33177 | 0.031673 | 0.26843 | 0.39512 | 0.23675 | 0.42679 | 9.547% | 19.09% | 28.64% | 0.31518 | 0.34836
U, ppm | 8.8288 | 0.49053 | 7.8477 | 9.8099 | 7.3572 | 10.3004 | 5.556% | 11.11% | 16.67% | 8.3873 | 9.2702
V, ppm | 200.6645 | 6.3019 | 188.0606 | 213.2684 | 181.7587 | 219.5703 | 3.141% | 6.281% | 9.422% | 190.6313 | 210.6977
W, ppm | 2.7063 | 0.33284 | 2.0406 | 3.372 | 1.7078 | 3.7048 | 12.3% | 24.6% | 36.9% | 2.571 | 2.8416
Y, ppm | 19.0781 | 1.0275 | 17.023 | 21.1331 | 15.9955 | 22.1607 | 5.386% | 10.77% | 16.16% | 18.1242 | 20.032
Yb, ppm | 2.2163 | 0.154 | 1.9083 | 2.5243 | 1.7543 | 2.6783 | 6.949% | 13.9% | 20.85% | 2.1055 | 2.3271
Zn, ppm | 26.6105 | 1.6348 | 23.3408 | 29.8801 | 21.706 | 31.5149 | 6.143% | 12.29% | 18.43% | 25.2799 | 27.941
Zr, ppm | 134.4601 | 5.8766 | 122.7068 | 146.2133 | 116.8302 | 152.09 | 4.371% | 8.741% | 13.11% | 127.7371 | 141.1831
Sulphuric Acid 5% Leach
Co, wt.% | 0.042001 | 0.0039951 | 0.034011 | 0.049991 | 0.030016 | 0.053986 | 9.512% | 19.02% | 28.54% | 0.039901 | 0.044101
Cu, wt.% | 0.83217 | 0.044266 | 0.74363 | 0.9207 | 0.69937 | 0.96496 | 5.319% | 10.64% | 15.96% | 0.79056 | 0.87378
Infrared Combustion
C, wt.% | 3.1219 | 0.080476 | 2.9609 | 3.2828 | 2.8805 | 3.3633 | 2.578% | 5.156% | 7.733% | 2.9658 | 3.278
S, wt.% | 1.1657 | 0.034006 | 1.0977 | 1.2337 | 1.0637 | 1.2677 | 2.917% | 5.835% | 8.752% | 1.1074 | 1.224
Borate Fusion XRF
Al2O3, wt.% | 11.4037 | 0.1693 | 11.0651 | 11.7423 | 10.8958 | 11.9116 | 1.485% | 2.969% | 4.454% | 10.8335 | 11.9738
BaO, ppm | 824.1849 | 71.5204 | 681.144 | 967.2258 | 609.6236 | 1038.7462 | 8.678% | 17.36% | 26.03% | 782.9756 | 865.3941
CaO, wt.% | 1.8724 | 0.02056 | 1.8313 | 1.9135 | 1.8107 | 1.9341 | 1.098% | 2.196% | 3.294% | 1.7788 | 1.966
Co, wt.% | 0.14455 | 0.0045617 | 0.13543 | 0.15367 | 0.13087 | 0.15824 | 3.156% | 6.312% | 9.467% | 0.13732 | 0.15178
Cu, wt.% | 1.8423 | 0.030284 | 1.7817 | 1.9029 | 1.7514 | 1.9331 | 1.644% | 3.288% | 4.932% | 1.7502 | 1.9344
Fe2O3, wt.% | 4.0116 | 0.070419 | 3.8707 | 4.1524 | 3.8003 | 4.2228 | 1.755% | 3.511% | 5.266% | 3.811 | 4.2121
K2O, wt.% | 2.4195 | 0.04084 | 2.3378 | 2.5012 | 2.297 | 2.542 | 1.688% | 3.376% | 5.064% | 2.2985 | 2.5405
MgO, wt.% | 7.1901 | 0.15593 | 6.8782 | 7.5019 | 6.7223 | 7.6579 | 2.169% | 4.337% | 6.506% | 6.8306 | 7.5496
MnO, wt.% | 0.025504 | 0.0033546 | 0.018794 | 0.032213 | 0.01544 | 0.035567 | 13.15% | 26.31% | 39.46% | 0.024228 | 0.026779
Ni, ppm | 56.6781 | 14.1055 | 28.467 | 84.8891 | 14.3615 | 98.9947 | 24.89% | 49.77% | 74.66% | 53.8442 | 59.512
P2O5, wt.% | 0.11534 | 0.0074856 | 0.10037 | 0.13031 | 0.092881 | 0.13779 | 6.49% | 12.98% | 19.47% | 0.10957 | 0.1211
SiO2, wt.% | 58.1847 | 0.5218 | 57.141 | 59.2283 | 56.6192 | 59.7501 | 0.8968% | 1.794% | 2.69% | 55.2754 | 61.0939
SO3, wt.% | 3.0023 | 0.10474 | 2.7928 | 3.2118 | 2.6881 | 3.3165 | 3.489% | 6.977% | 10.47% | 2.8522 | 3.1524
SrO, ppm | 102.1716 | 13.237 | 75.6977 | 128.6455 | 62.4607 | 141.8825 | 12.96% | 25.91% | 38.87% | 97.063 | 107.2802
TiO2, wt.% | 0.64592 | 0.015352 | 0.61522 | 0.67662 | 0.59986 | 0.69198 | 2.377% | 4.754% | 7.13% | 0.61362 | 0.67822
V2O5, ppm | 362.5762 | 67.3427 | 227.8908 | 497.2616 | 160.5481 | 564.6043 | 18.57% | 37.15% | 55.72% | 344.4474 | 380.705
Zr, ppm | 184.2244 | 30.5813 | 123.0618 | 245.3871 | 92.4805 | 275.9684 | 16.6% | 33.2% | 49.8% | 175.0132 | 193.4357
Thermogravimetry
LOI1000, wt.% | 9.9748 | 0.56274 | 8.8493 | 11.1003 | 8.2866 | 11.663 | 5.642% | 11.28% | 16.92% | 9.476 | 10.4735
Peroxide Fusion ICP
Al, wt.% | 5.8847 | 0.15687 | 5.571 | 6.1984 | 5.4141 | 6.3553 | 2.666% | 5.331% | 7.997% | 5.5905 | 6.1789
As, ppm | 59.9998 | 4.5967 | 50.8064 | 69.1932 | 46.2097 | 73.7899 | 7.661% | 15.32% | 22.98% | 56.9998 | 62.9998
B, ppm | 170.7019 | 19.5441 | 131.6137 | 209.7901 | 112.0696 | 229.3342 | 11.45% | 22.9% | 34.35% | 162.1668 | 179.237
Ba, ppm | 735.4886 | 17.6662 | 700.1563 | 770.8209 | 682.4901 | 788.4871 | 2.402% | 4.804% | 7.206% | 698.7142 | 772.263
Be, ppm | 3.76 | 0.51017 | 2.7397 | 4.7804 | 2.2295 | 5.2905 | 13.57% | 27.14% | 40.7% | 3.572 | 3.948
Bi, ppm | 2.1547 | 0.15138 | 1.8519 | 2.4574 | 1.7006 | 2.6088 | 7.025% | 14.05% | 21.08% | 2.047 | 2.2624
Ca, wt.% | 1.3528 | 0.070329 | 1.2121 | 1.4934 | 1.1418 | 1.5637 | 5.199% | 10.4% | 15.6% | 1.2851 | 1.4204
Ce, ppm | 126.4803 | 4.8418 | 116.7967 | 136.1639 | 111.9549 | 141.0057 | 3.828% | 7.656% | 11.48% | 120.1563 | 132.8043
Co, wt.% | 0.14509 | 0.0031721 | 0.13874 | 0.15143 | 0.13557 | 0.1546 | 2.186% | 4.373% | 6.559% | 0.13783 | 0.15234
Cr, ppm | 87.1448 | 16.959 | 53.2269 | 121.0628 | 36.2679 | 138.0218 | 19.46% | 38.92% | 58.38% | 82.7876 | 91.5021
Cs, ppm | 2.2573 | 0.1307 | 1.9959 | 2.5186 | 1.8652 | 2.6493 | 5.79% | 11.58% | 17.37% | 2.1444 | 2.3701
Cu, wt.% | 1.8305 | 0.047781 | 1.7349 | 1.926 | 1.6871 | 1.9738 | 2.61% | 5.221% | 7.831% | 1.739 | 1.922
Dy, ppm | 5.7834 | 0.32384 | 5.1358 | 6.4311 | 4.8119 | 6.755 | 5.599% | 11.2% | 16.8% | 5.4943 | 6.0726
Er, ppm | 3.1387 | 0.33806 | 2.4625 | 3.8148 | 2.1245 | 4.1528 | 10.77% | 21.54% | 32.31% | 2.9817 | 3.2956
Eu, ppm | 0.92213 | 0.059685 | 0.80276 | 1.0415 | 0.74307 | 1.1012 | 6.473% | 12.95% | 19.42% | 0.87602 | 0.96823
Fe, wt.% | 2.7424 | 0.089961 | 2.5624 | 2.9223 | 2.4725 | 3.0122 | 3.28% | 6.561% | 9.841% | 2.6052 | 2.8795
Ga, ppm | 23.0023 | 1.0202 | 20.9619 | 25.0428 | 19.9416 | 26.063 | 4.435% | 8.871% | 13.31% | 21.8522 | 24.1525
Gd, ppm | 6.1183 | 0.35724 | 5.4038 | 6.8328 | 5.0466 | 7.19 | 5.839% | 11.68% | 17.52% | 5.8124 | 6.4242
Ge, ppm | 2.01 | 0.11901 | 1.772 | 2.248 | 1.653 | 2.367 | 5.921% | 11.84% | 17.76% | 1.9095 | 2.1105
Ho, ppm | 1.1428 | 0.071306 | 1.0002 | 1.2854 | 0.92885 | 1.3567 | 6.24% | 12.48% | 18.72% | 1.0856 | 1.1999
In, ppm | 0.98272 | 0.10105 | 0.78063 | 1.1848 | 0.67958 | 1.2859 | 10.28% | 20.56% | 30.85% | 0.93359 | 1.0319
K, wt.% | 2.0317 | 0.063468 | 1.9047 | 2.1586 | 1.8413 | 2.2221 | 3.124% | 6.248% | 9.372% | 1.9301 | 2.1332
La, ppm | 65.9656 | 3.1615 | 59.6426 | 72.2887 | 56.481 | 75.4503 | 4.793% | 9.585% | 14.38% | 62.6674 | 69.2639
Li, ppm | 78.3646 | 5.3034 | 67.7579 | 88.9713 | 62.4545 | 94.2747 | 6.768% | 13.54% | 20.3% | 74.4464 | 82.2828
Lu, ppm | 0.47695 | 0.028077 | 0.4208 | 0.5331 | 0.39272 | 0.56118 | 5.887% | 11.77% | 17.66% | 0.4531 | 0.5008
Mg, wt.% | 4.3484 | 0.12624 | 4.0959 | 4.6008 | 3.9696 | 4.7271 | 2.903% | 5.806% | 8.709% | 4.1309 | 4.5658
Mn, wt.% | 0.019522 | 0.001962 | 0.015598 | 0.023446 | 0.013636 | 0.025408 | 10.05% | 20.1% | 30.15% | 0.018546 | 0.020498
Mo, ppm | 7.0692 | 0.79742 | 5.4744 | 8.664 | 4.677 | 9.4615 | 11.28% | 22.56% | 33.84% | 6.7158 | 7.4227
Nb, ppm | 19.4534 | 1.3092 | 16.835 | 22.0718 | 15.5258 | 23.381 | 6.73% | 13.46% | 20.19% | 18.4807 | 20.4261
Nd, ppm | 52.6152 | 2.0745 | 48.4663 | 56.7641 | 46.3918 | 58.8386 | 3.943% | 7.885% | 11.83% | 49.9844 | 55.246
Ni, ppm | 69.5639 | 8.8053 | 51.9533 | 87.1746 | 43.1479 | 95.9799 | 12.66% | 25.32% | 37.97% | 66.0857 | 73.0421
P, wt.% | 0.05178 | 0.0047156 | 0.042349 | 0.061212 | 0.037634 | 0.065927 | 9.107% | 18.21% | 27.32% | 0.049191 | 0.054369
Pb, ppm | 14.9633 | 1.2789 | 12.4056 | 17.5211 | 11.1267 | 18.7999 | 8.547% | 17.09% | 25.64% | 14.2152 | 15.7115
Pr, ppm | 14.4294 | 0.97323 | 12.4829 | 16.3758 | 11.5097 | 17.3491 | 6.745% | 13.49% | 20.23% | 13.7079 | 15.1509
Rb, ppm | 87.7846 | 3.5788 | 80.6271 | 94.9422 | 77.0483 | 98.521 | 4.077% | 8.154% | 12.23% | 83.3954 | 92.1739
Re, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 1.179 | 0.060318 | 1.0584 | 1.2996 | 0.99806 | 1.36 | 5.116% | 10.23% | 15.35% | 1.1201 | 1.238
Sb, ppm | 2.6653 | 0.20599 | 2.2533 | 3.0773 | 2.0473 | 3.2833 | 7.729% | 15.46% | 23.19% | 2.532 | 2.7985
Sc, ppm | 11.12 | 0.75922 | 9.6016 | 12.6384 | 8.8423 | 13.3977 | 6.827% | 13.65% | 20.48% | 10.564 | 11.676
Si, wt.% | 27.2817 | 0.85621 | 25.5693 | 28.9942 | 24.7131 | 29.8504 | 3.138% | 6.277% | 9.415% | 25.9176 | 28.6458
Sm, ppm | 7.6994 | 0.419 | 6.8614 | 8.5374 | 6.4425 | 8.9564 | 5.442% | 10.88% | 16.33% | 7.3145 | 8.0844
Sn, ppm | 3.0195 | 0.582 | 1.8555 | 4.1835 | 1.2735 | 4.7655 | 19.27% | 38.55% | 57.82% | 2.8685 | 3.1704
Sr, ppm | 78.2221 | 3.9518 | 70.3185 | 86.1257 | 66.3667 | 90.0776 | 5.052% | 10.1% | 15.16% | 74.311 | 82.1332
Ta, ppm | 1.5326 | 0.23738 | 1.0579 | 2.0074 | 0.82049 | 2.2448 | 15.49% | 30.98% | 46.47% | 1.456 | 1.6093
Tb, ppm | 0.96125 | 0.059802 | 0.84165 | 1.0809 | 0.78184 | 1.1407 | 6.221% | 12.44% | 18.66% | 0.91319 | 1.0093
Th, ppm | 13.6757 | 0.40099 | 12.8737 | 14.4777 | 12.4727 | 14.8786 | 2.932% | 5.864% | 8.796% | 12.9919 | 14.3595
Ti, wt.% | 0.37423 | 0.01246 | 0.34931 | 0.39915 | 0.33685 | 0.41161 | 3.33% | 6.659% | 9.989% | 0.35552 | 0.39294
Tl, ppm | 0.59784 | 0.066806 | 0.46422 | 0.73145 | 0.39742 | 0.79825 | 11.17% | 22.35% | 33.52% | 0.56794 | 0.62773
Tm, ppm | 0.48575 | 0.028655 | 0.42844 | 0.54306 | 0.39978 | 0.57172 | 5.899% | 11.8% | 17.7% | 0.46146 | 0.51004
U, ppm | 9.117 | 0.57217 | 7.9727 | 10.2614 | 7.4005 | 10.8336 | 6.276% | 12.55% | 18.83% | 8.6612 | 9.5729
V, ppm | 213.6536 | 10.6513 | 192.351 | 234.9563 | 181.6996 | 245.6076 | 4.985% | 9.971% | 14.96% | 202.9709 | 224.3363
W, ppm | 3.4021 | 0.55144 | 2.2992 | 4.505 | 1.7478 | 5.0564 | 16.21% | 32.42% | 48.63% | 3.232 | 3.5722
Y, ppm | 29.4088 | 1.0527 | 27.3035 | 31.5142 | 26.2508 | 32.5668 | 3.579% | 7.159% | 10.74% | 27.9384 | 30.8792
Yb, ppm | 3.1314 | 0.18788 | 2.7557 | 3.5072 | 2.5678 | 3.6951 | 6% | 12% | 18% | 2.9748 | 3.288
Aqua Regia Digestion
Ag, ppm | 0.23312 | 0.035193 | 0.16274 | 0.30351 | 0.12754 | 0.3387 | 15.1% | 30.19% | 45.29% | 0.22147 | 0.24478
Al, wt.% | 1.3109 | 0.11071 | 1.0895 | 1.5323 | 0.97879 | 1.643 | 8.445% | 16.89% | 25.34% | 1.2454 | 1.3765
As, ppm | 58.4178 | 2.9587 | 52.5004 | 64.3352 | 49.5418 | 67.2939 | 5.065% | 10.13% | 15.19% | 55.4969 | 61.3387
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 103.4893 | 14.1113 | 75.2668 | 131.7118 | 61.1556 | 145.8231 | 13.64% | 27.27% | 40.91% | 98.3149 | 108.6638
Be, ppm | 1.6285 | 0.13406 | 1.3604 | 1.8966 | 1.2263 | 2.0307 | 8.232% | 16.46% | 24.7% | 1.5471 | 1.7099
Bi, ppm | 1.9093 | 0.093359 | 1.7226 | 2.096 | 1.6293 | 2.1894 | 4.89% | 9.779% | 14.67% | 1.8139 | 2.0048
Ca, wt.% | 1.3184 | 0.048048 | 1.2223 | 1.4145 | 1.1742 | 1.4625 | 3.645% | 7.289% | 10.93% | 1.2524 | 1.3843
Cd, ppm | 0.19621 | 0.013458 | 0.16929 | 0.22313 | 0.15583 | 0.23658 | 6.859% | 13.72% | 20.58% | 0.1864 | 0.20602
Ce, ppm | 43.3083 | 5.0599 | 33.1885 | 53.4281 | 28.1286 | 58.488 | 11.68% | 23.37% | 35.05% | 41.1429 | 45.4737
Co, wt.% | 0.13079 | 0.011681 | 0.10743 | 0.15415 | 0.095746 | 0.16583 | 8.931% | 17.86% | 26.79% | 0.12425 | 0.13733
Cr, ppm | 28.4997 | 1.9817 | 24.5364 | 32.463 | 22.5547 | 34.4447 | 6.953% | 13.91% | 20.86% | 27.0747 | 29.9247
Cs, ppm | 0.4544 | 0.071177 | 0.31205 | 0.59675 | 0.24087 | 0.66793 | 15.66% | 31.33% | 46.99% | 0.43168 | 0.47712
Cu, wt.% | 1.8171 | 0.03316 | 1.7508 | 1.8834 | 1.7176 | 1.9166 | 1.825% | 3.65% | 5.475% | 1.7263 | 1.908
Dy, ppm | 1.1679 | 0.11609 | 0.9357 | 1.4 | 0.81961 | 1.5161 | 9.94% | 19.88% | 29.82% | 1.1095 | 1.2263
Er, ppm | 0.51885 | 0.038429 | 0.44199 | 0.5957 | 0.40356 | 0.63413 | 7.407% | 14.81% | 22.22% | 0.4929 | 0.54479
Eu, ppm | 0.39055 | 0.048866 | 0.29282 | 0.48828 | 0.24395 | 0.53715 | 12.51% | 25.02% | 37.54% | 0.37102 | 0.41008
Fe, wt.% | 2.4576 | 0.070042 | 2.3175 | 2.5977 | 2.2475 | 2.6678 | 2.85% | 5.7% | 8.55% | 2.3347 | 2.5805
Ga, ppm | 7.3108 | 0.73907 | 5.8327 | 8.789 | 5.0936 | 9.528 | 10.11% | 20.22% | 30.33% | 6.9453 | 7.6764
Gd, ppm | 2.1919 | 0.17403 | 1.8438 | 2.5399 | 1.6698 | 2.714 | 7.94% | 15.88% | 23.82% | 2.0823 | 2.3015
Hf, ppm | 0.32104 | 0.034313 | 0.25241 | 0.38966 | 0.2181 | 0.42398 | 10.69% | 21.38% | 32.06% | 0.30499 | 0.33709
Hg, ppm | 0.047214 | 0.0095351 | 0.028144 | 0.066284 | 0.018609 | 0.075819 | 20.2% | 40.39% | 60.59% | 0.044854 | 0.049575
Ho, ppm | 0.19551 | 0.016449 | 0.16261 | 0.22841 | 0.14616 | 0.24486 | 8.413% | 16.83% | 25.24% | 0.18573 | 0.20528
In, ppm | 0.62623 | 0.034923 | 0.55639 | 0.69608 | 0.52146 | 0.731 | 5.577% | 11.15% | 16.73% | 0.59492 | 0.65754
K, wt.% | 0.20498 | 0.020633 | 0.16372 | 0.24625 | 0.14309 | 0.26688 | 10.07% | 20.13% | 30.2% | 0.19473 | 0.21523
La, ppm | 19.2317 | 2.0279 | 15.1759 | 23.2874 | 13.148 | 25.3153 | 10.54% | 21.09% | 31.63% | 18.2701 | 20.1932
Li, ppm | 49.8796 | 6.5919 | 36.6959 | 63.0634 | 30.104 | 69.6553 | 13.22% | 26.43% | 39.65% | 47.3857 | 52.3736
Lu, ppm | 0.065362 | 0.0057893 | 0.053783 | 0.076941 | 0.047994 | 0.08273 | 8.857% | 17.71% | 26.57% | 0.062094 | 0.06863
Mg, wt.% | 3.5321 | 0.16684 | 3.1985 | 3.8658 | 3.0316 | 4.0327 | 4.723% | 9.447% | 14.17% | 3.3555 | 3.7088
Mn, wt.% | 0.017955 | 0.0008442 | 0.016266 | 0.019643 | 0.015422 | 0.020487 | 4.702% | 9.404% | 14.11% | 0.017057 | 0.018852
Mo, ppm | 6.5632 | 0.2846 | 5.994 | 7.1324 | 5.7094 | 7.417 | 4.336% | 8.673% | 13.01% | 6.2351 | 6.8914
Na, wt.% | 0.0095787 | 0.001333 | 0.0069127 | 0.012245 | 0.0055797 | 0.013578 | 13.92% | 27.83% | 41.75% | 0.0090997 | 0.010058
Nd, ppm | 17.4893 | 3.142 | 11.2053 | 23.7733 | 8.0633 | 26.9153 | 17.97% | 35.93% | 53.9% | 16.6148 | 18.3638
Ni, ppm | 60.0636 | 2.2928 | 55.478 | 64.6491 | 53.1852 | 66.9419 | 3.817% | 7.635% | 11.45% | 57.0604 | 63.0667
P, wt.% | 0.039341 | 0.0023351 | 0.034671 | 0.044012 | 0.032336 | 0.046347 | 5.936% | 11.87% | 17.81% | 0.037374 | 0.041309
Pb, ppm | 9.1344 | 1.3597 | 6.4149 | 11.8538 | 5.0552 | 13.2135 | 14.89% | 29.77% | 44.66% | 8.6776 | 9.5911
Pr, ppm | 4.6782 | 0.81498 | 3.0482 | 6.3082 | 2.2333 | 7.1232 | 17.42% | 34.84% | 52.26% | 4.4443 | 4.9121
Rb, ppm | 8.7004 | 1.1754 | 6.3497 | 11.0511 | 5.1743 | 12.2265 | 13.51% | 27.02% | 40.53% | 8.2654 | 9.1354
Re, ppm | 0.011667 | 0.0010361 | 0.0095945 | 0.013739 | 0.0085584 | 0.014775 | 8.881% | 17.76% | 26.64% | 0.011083 | 0.01225
S, wt.% | 1.1612 | 0.05407 | 1.053 | 1.2693 | 0.99898 | 1.3234 | 4.656% | 9.313% | 13.97% | 1.1031 | 1.2192
Sb, ppm | 1.6192 | 0.29021 | 1.0388 | 2.1996 | 0.74857 | 2.4898 | 17.92% | 35.85% | 53.77% | 1.5382 | 1.7001
Sc, ppm | 3.6846 | 0.26687 | 3.1509 | 4.2184 | 2.884 | 4.4852 | 7.243% | 14.49% | 21.73% | 3.5004 | 3.8689
Se, ppm | 3.5038 | 0.40136 | 2.7011 | 4.3065 | 2.2997 | 4.7079 | 11.45% | 22.91% | 34.36% | 3.3286 | 3.679
Sm, ppm | 3.0035 | 0.46441 | 2.0747 | 3.9323 | 1.6103 | 4.3967 | 15.46% | 30.92% | 46.39% | 2.8533 | 3.1537
Sn, ppm | 0.83287 | 0.045057 | 0.74276 | 0.92299 | 0.6977 | 0.96804 | 5.41% | 10.82% | 16.23% | 0.79123 | 0.87452
Sr, ppm | 29.6478 | 2.468 | 24.7119 | 34.5838 | 22.244 | 37.0517 | 8.324% | 16.65% | 24.97% | 28.1654 | 31.1302
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.25977 | 0.022942 | 0.21389 | 0.30566 | 0.19095 | 0.3286 | 8.832% | 17.66% | 26.49% | 0.24678 | 0.27276
Te, ppm | 0.048338 | 0.01569 | 0.016957 | 0.079718 | 0.001267 | 0.095409 | 32.46% | 64.92% | 97.38% | 0.045921 | 0.050755
Th, ppm | 6.1023 | 0.57387 | 4.9545 | 7.25 | 4.3807 | 7.8239 | 9.404% | 18.81% | 28.21% | 5.7972 | 6.4074
U, ppm | 4.6116 | 0.29016 | 4.0313 | 5.1919 | 3.7411 | 5.4821 | 6.292% | 12.58% | 18.88% | 4.381 | 4.8422
V, ppm | 45.9389 | 5.8534 | 34.2321 | 57.6457 | 28.3787 | 63.4991 | 12.74% | 25.48% | 38.23% | 43.6419 | 48.2358
W, ppm | 1.0022 | 0.18514 | 0.63187 | 1.3724 | 0.44672 | 1.5576 | 18.47% | 36.95% | 55.42% | 0.95205 | 1.0523
Y, ppm | 4.5461 | 0.30489 | 3.9363 | 5.1559 | 3.6314 | 5.4608 | 6.707% | 13.41% | 20.12% | 4.3188 | 4.7734
Yb, ppm | 0.42848 | 0.038502 | 0.35147 | 0.50548 | 0.31297 | 0.54398 | 8.986% | 17.97% | 26.96% | 0.40705 | 0.4499
Zn, ppm | 22.872 | 1.53 | 19.8121 | 25.9319 | 18.2821 | 27.4619 | 6.689% | 13.38% | 20.07% | 21.7284 | 24.0156
Zr, ppm | 10.9703 | 1.3174 | 8.3355 | 13.6051 | 7.0181 | 14.9225 | 12.01% | 24.02% | 36.03% | 10.4218 | 11.5188
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 550b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AL-5S*AAS | 5% sulphuric acid laech with atomic absorption spectroscopy
AL-5S*OES | 5% sulphuric acid leach with inductively coupled plasma optical emission spectroscopy
AL-5S*OES/AAS | 5% sulphuric acid leach with AAS or ICP-OES finish
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Text Values:
NA | Unable to report due to QC failure (Lab 19)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 550b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Labwest Minerals Analysis, Perth, WA, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 550b (Certified Value 0.323 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.44 | 0.252 | 0.5 | 0.3 | 1.2 | 0.27 | 0.38789 | 0.35 | 0.31 | 0.63 | 0.36 | 0.28 | 0.34356 | 0.4 | 0.34 | 0.32 | < 0.3 | 0.5 | 0.3 | 0.3 | 0.28 |  |  | 1
| 1 | 2 | 0.42 | 0.278 | 0.5 | 0.3 | 1.1 | 0.3 | 0.31508 | 0.39 | 0.34 | 0.3 | 0.39 | 0.3 | 0.39747 | 0.4 | 0.37 | 0.31 | 0.4 | 0.3 | 0.3 | 0.29 | 0.23 |  |  | 4
| 1 | 3 | 0.36 | 0.454 | 0.5 | 0.3 | 1.2 | 0.28 | 0.26277 | 0.26 | 0.34 | 0.31 | 0.37 | 0.46 | 0.34173 | 0.4 | 0.3 | 0.34 | 0.3 | 0.3 | 0.3 | 0.31 | 0.3 |  |  | 16
| 1 | 4 | 0.29 | 0.264 | 0.5 | 0.3 | 1.2 | 0.29 | 0.37018 | 0.34 | 0.48 | 0.29 | 0.31 | 0.62 | 0.34705 | 0.4 | 0.29 | 0.36 | < 0.3 | 0.2 | 0.2 | 0.35 | 0.28 |  |  | 0.32344
| 1 | 5 | 0.39 | 0.231 | 0.4 | 0.3 | 1.1 | 0.27 | 0.26068 | 0.4 | 0.26 | 0.68 | 0.29 | 0.38 | 0.3842 | 0.4 | 0.23 | 0.33 | < 0.3 | 0.3 | 0.3 | 0.32 | 0.45 |  |  | 7
| 1 | 6 | 0.29 | 0.236 | 0.5 | 0.3 | 1.2 | 0.3 | 0.31156 | 0.34 | 0.29 | 0.33 | 0.37 | 0.27 | 0.36663 | 0.4 | 0.34 | 0.33 | < 0.3 | 0.3 | 0.4 | 0.32 | 0.39
| Mean |  | 0.365 | 0.28583 | 0.48333 | 0.3 | 1.1667 | 0.285 | 0.31803 | 0.34667 | 0.33667 | 0.42333 | 0.34833 | 0.385 | 0.36344 | 0.4 | 0.31167 | 0.33167 | 0.35 | 0.31667 | 0.3 | 0.315 | 0.32167
| Median |  | 0.375 | 0.258 | 0.5 | 0.3 | 1.2 | 0.285 | 0.31332 | 0.345 | 0.325 | 0.32 | 0.365 | 0.34 | 0.35684 | 0.4 | 0.32 | 0.33 | 0.35 | 0.3 | 0.3 | 0.315 | 0.29
| Std Dev. |  | 0.064109 | 0.084206 | 0.040825 | 0 | 0.05164 | 0.013784 | 0.052901 | 0.049666 | 0.076594 | 0.18063 | 0.0392 | 0.1359 | 0.023386 | 6.0809e-17 | 0.049565 | 0.017224 | 0.070711 | 0.098319 | 0.063246 | 0.020736 | 0.081833
| Rel.Std.Dev. |  | 17.56% | 29.46% | 8.447% | 0% | 4.426% | 4.837% | 16.63% | 14.33% | 22.75% | 42.67% | 11.25% | 35.3% | 6.435% | 1.52e-14% | 15.9% | 5.193% | 20.2% | 31.05% | 21.08% | 6.583% | 25.44%
| PDM3 |  | 12.85% | -11.63% | 49.44% | -7.246% | 260.7% | -11.88% | -1.672% | 7.182% | 4.091% | 30.89% | 7.698% | 19.03% | 12.37% | 23.67% | -3.639% | 2.545% | 8.213% | -2.093% | -7.246% | -2.608% | -0.5471%
| Z-Score (Absolute) |  | 1 | 1.56 | NA | NA | NA | 1.59 | 0.52 | 0.4 | 0.08 | 2.89 | 0.46 | 1.65 | 0.95 | NA | 0.73 | 0.08 | NA | NA | NA | 0.62 | 0.4
Results from laboratories 3, 4, 7, 16, 19, 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 550b (Certified Value 5.88 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.8174 | 6.17 | 5.63 | 5.86 | 5.8412 | 5.43 | 5.88 | 5.67 | 6.0618 | 5.91 | 5.8 | 5.76 | 5.75 | 5.84 | 5.8778 | 5.78 | 5.7 | 6.0184 | 5.5307 | 5.87 | 6.16 | 5.79 | 5.05 | 1
| 1 | 2 | 5.8746 | 6.22 | 5.63 | 5.9 | 5.8452 | 5.76 | 5.91 | 5.62 | 5.9946 | 5.96 | 5.76 | 5.73 | 5.74 | 5.9 | 5.8391 | 5.85 | 5.71 | 6.0867 | 5.9243 | 5.96 | 6.24 | 5.77 | 5.19 | #N/A
| 1 | 3 | 5.9623 | 6.22 | 5.67 | 5.85 | 5.8364 | 5.88 | 5.88 | 5.71 | 5.9508 | 5.84 | 5.75 | 5.41 | 6.2 | 5.91 | 5.8746 | 5.91 | 5.74 | 6.0057 | 5.9573 | 5.84 | 6.21 | 5.87 | 5.21 | 16
| 1 | 4 | 5.9551 | 6.22 | 5.59 | 5.92 | 5.8319 | 5.83 | 5.93 | 5.63 | 5.9571 | 5.89 | 5.8 | 5.41 | 6.1 | 5.86 | 5.9441 | 5.84 | 5.76 | 6.0739 | 5.8245 | 5.82 | 6.31 | 6.08 | 5.18 | 5.8761
| 1 | 5 | 5.7603 | 6.24 | 5.57 | 5.92 | 5.8586 | 6.25 | 5.85 | 5.65 | 6.0842 | 5.77 | 5.71 | 5.51 | 6.18 | 5.93 | 5.7968 | 5.96 | 5.83 | 5.8741 | 5.7678 | 5.9 | 6.31 | 6.02 | 5.25 | 8
| 1 | 6 | 6.1628 | 6.06 | 5.65 | 5.96 | 5.8462 | 5.35 | 5.94 | 5.69 | 5.9837 | 5.96 | 5.39 | 5.63 | 6.2 | 5.9 | 5.9242 | 5.82 | 5.74 | 6.0558 | 5.8265 | 6.04 | 6.19 | 5.78 | 5.24
| Mean |  | 5.9221 | 6.1883 | 5.6233 | 5.9017 | 5.8432 | 5.75 | 5.8983 | 5.6617 | 6.0054 | 5.8883 | 5.7017 | 5.575 | 6.0283 | 5.89 | 5.8761 | 5.86 | 5.7467 | 6.0191 | 5.8052 | 5.905 | 6.2367 | 5.885 | 5.1867
| Median |  | 5.9148 | 6.22 | 5.63 | 5.91 | 5.8432 | 5.795 | 5.895 | 5.66 | 5.9892 | 5.9 | 5.755 | 5.57 | 6.14 | 5.9 | 5.8762 | 5.845 | 5.74 | 6.0371 | 5.8255 | 5.885 | 6.225 | 5.83 | 5.2
| Std Dev. |  | 0.14145 | 0.067057 | 0.037238 | 0.041191 | 0.0092619 | 0.32735 | 0.034303 | 0.034881 | 0.055282 | 0.073598 | 0.15639 | 0.15489 | 0.22257 | 0.033466 | 0.054044 | 0.064807 | 0.046332 | 0.077621 | 0.15159 | 0.082401 | 0.062503 | 0.13398 | 0.072296
| Rel.Std.Dev. |  | 2.389% | 1.084% | 0.6622% | 0.6979% | 0.1585% | 5.693% | 0.5816% | 0.6161% | 0.9206% | 1.25% | 2.743% | 2.778% | 3.692% | 0.5682% | 0.9197% | 1.106% | 0.8062% | 1.29% | 2.611% | 1.395% | 1.002% | 2.277% | 1.394%
| PDM3 |  | 0.7832% | 5.314% | -4.301% | 0.4357% | -0.5588% | -2.145% | 0.379% | -3.649% | 2.2% | 0.2088% | -2.968% | -5.124% | 2.591% | 0.2372% | 0.0002903% | -0.2734% | -2.202% | 2.434% | -1.206% | 0.4925% | 6.137% | 0.1521% | -11.73%
| Z-Score (Absolute) |  | 0.21 | 1.7 | 1.47 | 0.09 | 0.23 | 0.76 | 0.07 | 1.25 | 0.67 | 0.02 | 1.03 | 1.74 | 0.8 | 0.03 | 0.05 | 0.14 | 0.77 | 0.75 | 0.45 | 0.11 | 1.97 | 0 | 3.91
| Analytical results for As in OREAS 550b (Certified Value 61 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
1
| 1 | 1 | 59.5 | 55.7 | 71.5 | 59 | 61 | 62 | 60.127 | 58.8 | 66.7 | 62.2 | 59.9 | 59.9 | 63.2365 | 57.05 | 67 | 61.8 | 60.6 | 53 | 52.6 | 56.6 | 57 |  |  | 1
| 1 | 2 | 60.6 | 54.2 | 66.8 | 58 | 61 | 60 | 62.4326 | 60.5 | 65.4 | 62.8 | 58.2 | 61.3 | 61.86 | 59.5 | 63 | 65.1 | 62.9 | 55 | 51.8 | 56.1 | 64 |  |  | 5
| 1 | 3 | 59.8 | 57.3 | 68.8 | 57 | 61 | 61 | 58.5137 | 59.6 | 66 | 62.2 | 62.4 | 63.4 | 62.5369 | 59.45 | 63 | 61.3 | 63.8 | 57 | 53.3 | 55.4 | 68 |  |  | 16
| 1 | 4 | 59.4 | 57.1 | 66.9 | 58 | 62 | 60 | 62.2617 | 58.3 | 65.3 | 62.4 | 61.7 | 61.7 | 63.5472 | 58.85 | 61 | 62.3 | 63.3 | 52 | 51.5 | 54 | 64 |  |  | 60.5426
| 1 | 5 | 60.4 | 57.3 | 66 | 59 | 60 | 63 | 60.4821 | 59.9 | 65.6 | 63.9 | 61.1 | 61.8 | 64.288 | 60.03 | 62 | 61.1 | 62.7 | 53 | 52.8 | 54.6 | 65 |  |  | 9
| 1 | 6 | 59.7 | 54.9 | 69.1 | 58 | 61 | 59 | 59.575 | 61.3 | 65.4 | 62.1 | 60.8 | 60.9 | 64.9971 | 59.72 | 59 | 64.2 | 63.5 | 57 | 55.7 | 56.7 | 60
| Mean |  | 59.9 | 56.0833 | 68.1833 | 58.1667 | 61 | 60.8333 | 60.5653 | 59.7333 | 65.7333 | 62.6 | 60.6833 | 61.5 | 63.4109 | 59.1 | 62.5 | 62.6333 | 62.8 | 54.5 | 52.95 | 55.5667 | 63
| Median |  | 59.75 | 56.4 | 67.85 | 58 | 61 | 60.5 | 60.3046 | 59.75 | 65.5 | 62.3 | 60.95 | 61.5 | 63.3918 | 59.475 | 62.5 | 62.05 | 63.1 | 54 | 52.7 | 55.75 | 64
| Std Dev. |  | 0.4899 | 1.3482 | 2.0272 | 0.75277 | 0.63246 | 1.472 | 1.5332 | 1.0967 | 0.53541 | 0.68411 | 1.4798 | 1.1576 | 1.1405 | 1.0767 | 2.6646 | 1.6415 | 1.1489 | 2.1679 | 1.5003 | 1.1003 | 3.8987
| Rel.Std.Dev. |  | 0.8179% | 2.404% | 2.973% | 1.294% | 1.037% | 2.42% | 2.531% | 1.836% | 0.8145% | 1.093% | 2.438% | 1.882% | 1.799% | 1.822% | 4.263% | 2.621% | 1.829% | 3.978% | 2.833% | 1.98% | 6.188%
| PDM3 |  | -1.061% | -7.366% | 12.62% | -3.924% | 0.7555% | 0.4802% | 0.03756% | -1.337% | 8.574% | 3.398% | 0.2324% | 1.581% | 4.738% | -2.383% | 3.233% | 3.453% | 3.729% | -9.981% | -12.54% | -8.219% | 4.059%
| Z-Score (Absolute) |  | 0.35 | 1.78 | 2.75 | 1 | 0.06 | 0 | 0.1 | 0.41 | 1.84 | 0.66 | 0.06 | 0.25 | 0.97 | 0.65 | 0.62 | 0.67 | 0.74 | 2.37 | 2.95 | 1.97 | 0.81
| Analytical results for B in OREAS 550b (Indicative Value 18.5 ppm) | [24 empty] | Indicative
B | Round | Replicate | Lab | [22 empty] | 1
| No | No | 07 | [22 empty] | ppm
|  |  | 4A*OES | [22 empty] | 1
1
| 1 | 1 | 19 | [22 empty] | 1
| 1 | 2 | 19 | [22 empty] | 1
| 1 | 3 | 18 | [22 empty] | 16
| 1 | 4 | 19 | [22 empty] | 18.5
| 1 | 5 | 18 | [22 empty] | 7
| 1 | 6 | 18
| Mean |  | 18.5
| Median |  | 18.5
| Std Dev. |  | 0.54772
| Rel.Std.Dev. |  | 2.961%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 550b (Certified Value 735 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
0
| 1 | 1 | 783 | 426 | 750 | 755 | 647 | 785 | 763 | 705.9798 | 730 | 740 | 720 | 750 | 740 | 679.7808 | 708 | 787.5 | 767 | 671 | 755 | 710 | 614 |  |  | 1
| 1 | 2 | 789 | 324 | 717 | 751 | 701 | 810 | 750 | 700.0195 | 750 | 730 | 730 | 740 | 750 | 683.2443 | 708 | 791.2 | 719 | 445 | 766 | 723 | 644 |  |  | #N/A
| 1 | 3 | 777.4 | 394 | 718 | 745 | 718 | 772 | 773 | 704.0923 | 730 | 740 | 690 | 800 | 750 | 692.2769 | 714 | 786.3 | 714 | 314 | 765 | 702 | 656 |  |  | 16
| 1 | 4 | 858.7 | 193 | 736 | 754 | 746 | 786 | 757 | 701.9323 | 730 | 740 | 690 | 790 | 740 | 700.7342 | 713 | 802 | 701 | 377 | 763 | 706 | 643 |  |  | 735.3056
| 1 | 5 | 739.8 | 327 | 743 | 749 | 840 | 793 | 751 | 706.3345 | 720 | 730 | 700 | 790 | 750 | 679.189 | 722 | 767.2 | 699 | 353 | 772 | 728 | 649 |  |  | 10
| 1 | 6 | 785.5 | 336 | 736 | 754 | 505 | 805 | 774 | 697.9174 | 750 | 700 | 710 | 790 | 750 | 707.8946 | 715 | 794.2 | 696 | 416 | 768 | 745 | 623
| Mean |  | 788.9 | 333.3333 | 733.3333 | 751.3333 | 692.8333 | 791.8333 | 761.3333 | 702.7126 | 735 | 730 | 706.6667 | 776.6667 | 746.6667 | 690.52 | 713.3333 | 788.0667 | 716 | 429.3333 | 764.8333 | 719 | 638.1667
| Median |  | 784.25 | 331.5 | 736 | 752.5 | 709.5 | 789.5 | 760 | 703.0123 | 730 | 735 | 705 | 790 | 750 | 687.7606 | 713.5 | 789.35 | 707.5 | 396.5 | 765.5 | 716.5 | 643.5
| Std Dev. |  | 38.6302 | 80.1191 | 13.3217 | 3.8297 | 111.8632 | 13.9917 | 10.5198 | 3.3631 | 12.2474 | 15.4919 | 16.3299 | 25.0333 | 5.164 | 11.8824 | 5.2026 | 11.6656 | 26.5631 | 127.0381 | 5.7067 | 16.1988 | 16.1668
| Rel.Std.Dev. |  | 4.897% | 24.04% | 1.817% | 0.5097% | 16.15% | 1.767% | 1.382% | 0.4786% | 1.666% | 2.122% | 2.311% | 3.223% | 0.6916% | 1.721% | 0.7293% | 1.48% | 3.71% | 29.59% | 0.7461% | 2.253% | 2.533%
| PDM3 |  | 7.289% | -54.67% | -0.2682% | 2.18% | -5.776% | 7.688% | 3.54% | -4.433% | -0.04156% | -0.7215% | -3.895% | 5.625% | 1.545% | -6.091% | -2.988% | 7.175% | -2.626% | -41.61% | 4.016% | -2.218% | -13.21%
| Z-Score (Absolute) |  | 1.27 | 8.54 | 0.07 | 0.46 | 0.8 | 1.33 | 0.67 | 0.59 | 0.11 | 0 | 0.5 | 1 | 0.36 | 0.85 | 0.36 | 1.25 | 0.3 | 6.47 | 0.75 | 0.24 | 1.98
| Analytical results for Be in OREAS 550b (Certified Value 3.34 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 3.53 | 4 | 4.3 | 3.2 | 7.6362 | 3.21 | 2.9 | 3.5 | 3.6144 | 3.31 | 3.34 | 3.29 | 3.31 | 3.33 | 3.3543 | 3.65 | 3.72 | 3.28 | 2.81 | 3 | 3 | 3.27 | 2.77 | 1
| 1 | 2 | 3.6 | 5 | 3.9 | 3.1 | 7.6693 | 3.1 | 3 | 3.4 | 3.7103 | 3.33 | 3.28 | 3.36 | 3.19 | 3.32 | 3.4664 | 3.55 | 3.63 | 3.45 | 2.96 | 4 | 3 | 3.27 | 2.85 | 6
| 1 | 3 | 3.72 | 5 | 4.1 | 3.1 | 7.6094 | 3.44 | 2.6 | 3.3 | 3.6129 | 3.29 | 3.28 | 3.17 | 3.37 | 3.32 | 3.518 | 3.4 | 3.6 | 3.4 | 3 | 4 | 3 | 3.57 | 2.7 | 16
| 1 | 4 | 3.6 | 4 | 4.1 | 3.1 | 7.7278 | 3.24 | 3 | 3.6 | 3.7171 | 3.33 | 3.32 | 3.16 | 3.34 | 3.32 | 3.396 | 3.47 | 3.53 | 3.39 | 2.92 | 3 | 3 | 3.1 | 2.85 | 3.3393
| 1 | 5 | 3.56 | 4 | 4.1 | 3.3 | 7.6248 | 3.34 | 2.9 | 3.4 | 3.7004 | 3.24 | 3.26 | 3.22 | 3.35 | 3.37 | 3.4624 | 3.4 | 3.74 | 3.37 | 2.92 | 3 | 3 | 3.42 | 2.78 | 11
| 1 | 6 | 3.39 | 4 | 4.1 | 3.3 | 7.668 | 3.22 | 2.9 | 3.5 | 3.6039 | 3.31 | 3.28 | 3.27 | 3.37 | 3.36 | 3.4684 | 3.67 | 3.56 | 3.46 | 2.91 | 4 | 3 | 3.18 | 2.7
| Mean |  | 3.5667 | 4.3333 | 4.1 | 3.1833 | 7.6559 | 3.2583 | 2.8833 | 3.45 | 3.6598 | 3.3017 | 3.2933 | 3.245 | 3.3217 | 3.3367 | 3.4442 | 3.5233 | 3.63 | 3.3917 | 2.92 | 3.5 | 3 | 3.3017 | 2.775
| Median |  | 3.58 | 4 | 4.1 | 3.15 | 7.6521 | 3.23 | 2.9 | 3.45 | 3.6574 | 3.31 | 3.28 | 3.245 | 3.345 | 3.325 | 3.4644 | 3.51 | 3.615 | 3.395 | 2.92 | 3.5 | 3 | 3.27 | 2.775
| Std Dev. |  | 0.10801 | 0.5164 | 0.12649 | 0.098319 | 0.042457 | 0.11737 | 0.1472 | 0.10488 | 0.054531 | 0.033714 | 0.030111 | 0.076616 | 0.06824 | 0.022509 | 0.058748 | 0.11961 | 0.084853 | 0.064936 | 0.063561 | 0.54772 | 0 | 0.1694 | 0.067157
| Rel.Std.Dev. |  | 3.028% | 11.92% | 3.085% | 3.089% | 0.5546% | 3.602% | 5.105% | 3.04% | 1.49% | 1.021% | 0.9143% | 2.361% | 2.054% | 0.6746% | 1.706% | 3.395% | 2.338% | 1.915% | 2.177% | 15.65% | 0% | 5.131% | 2.42%
| PDM3 |  | 6.81% | 29.77% | 22.78% | -4.669% | 129.3% | -2.423% | -13.65% | 3.316% | 9.6% | -1.126% | -1.375% | -2.823% | -0.5268% | -0.07762% | 3.144% | 5.512% | 8.707% | 1.569% | -12.56% | 4.814% | -10.16% | -1.126% | -16.9%
| Z-Score (Absolute) |  | 1.2 | NA | 3.9 | 0.74 | 21.88 | 0.36 | 2.25 | 0.61 | 1.67 | 0.14 | 0.18 | 0.43 | 0.04 | 0.04 | 0.58 | 0.98 | 1.52 | 0.32 | 2.07 | NA | NA | 0.14 | 2.8
Results from laboratories 2, 20 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 550b (Certified Value 2.1 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 2.28 | 1.98 | 2.75 | 2.09 | 2.21 | 1.89 | 2.2139 | 1.94 | 2.25 | 2.04 | 2.24 | 2.09 | 2.1622 | 2.08 | 2.12 | 2.35 | NA | 2.1 | 2.1 | 1.95 | < 5 |  |  | 1
| 1 | 2 | 2.21 | 2.04 | 2.11 | 2.07 | 2.32 | 1.99 | 2.2646 | 2 | 2.16 | 2.01 | 1.91 | 2.1 | 2.1405 | 2.11 | 2.24 | 2.38 | NA | 2 | 2.2 | 1.97 | < 5 |  |  | 7
| 1 | 3 | 2.26 | 1.92 | 2.06 | 2.06 | 2.37 | 1.88 | 2.2329 | 1.98 | 2.13 | 2 | 2.05 | 2.11 | 2.0793 | 2.09 | 2.14 | 2.24 | NA | 2.1 | 2.2 | 1.98 | < 5 |  |  | 16
| 1 | 4 | 2.21 | 1.84 | 2.07 | 2.04 | 2.34 | 1.92 | 2.2795 | 1.96 | 2.18 | 2.03 | 2.1 | 2.12 | 2.1837 | 2.13 | 2.1 | 2.29 | NA | 2 | 2.1 | 1.92 | < 5 |  |  | 2.1024
| 1 | 5 | 2.17 | 1.9 | 2 | 2.02 | 2.19 | 1.98 | 2.1926 | 1.97 | 2.13 | 1.98 | 2.06 | 2.1 | 2.1118 | 2.09 | 2.17 | 2.2 | NA | 2.1 | 2.2 | 2 | < 5 |  |  | 12
| 1 | 6 | 2.25 | 1.72 | 2.01 | 2.04 | 2.35 | 1.9 | 2.2671 | 2.06 | 2.11 | 2.02 | 2.05 | 2.1 | 2.143 | 2.05 | 2.09 | 2.26 | NA | 2.1 | 2.3 | 2.02 | < 5
| Mean |  | 2.23 | 1.9 | 2.1667 | 2.0533 | 2.2967 | 1.9267 | 2.2418 | 1.985 | 2.16 | 2.0133 | 2.0683 | 2.1033 | 2.1368 | 2.0917 | 2.1433 | 2.2867 |  | 2.0667 | 2.1833 | 1.9733
| Median |  | 2.23 | 1.91 | 2.065 | 2.05 | 2.33 | 1.91 | 2.2487 | 1.975 | 2.145 | 2.015 | 2.055 | 2.1 | 2.1418 | 2.09 | 2.13 | 2.275 |  | 2.1 | 2.2 | 1.975
| Std Dev. |  | 0.040497 | 0.11171 | 0.28863 | 0.025033 | 0.076855 | 0.047188 | 0.034243 | 0.041833 | 0.050596 | 0.021602 | 0.1061 | 0.010328 | 0.036935 | 0.027142 | 0.055377 | 0.06802 |  | 0.05164 | 0.075277 | 0.03559
| Rel.Std.Dev. |  | 1.816% | 5.88% | 13.32% | 1.219% | 3.346% | 2.449% | 1.528% | 2.107% | 2.342% | 1.073% | 5.13% | 0.491% | 1.729% | 1.298% | 2.584% | 2.975% |  | 2.499% | 3.448% | 1.804%
| PDM3 |  | 6.068% | -9.628% | 3.055% | -2.335% | 9.239% | -8.36% | 6.628% | -5.586% | 2.738% | -4.238% | -1.622% | 0.04287% | 1.633% | -0.512% | 1.945% | 8.763% |  | -1.701% | 3.848% | -6.14%
| Z-Score (Absolute) |  | 0.85 | 1.71 | 0.36 | 0.52 | 1.37 | 1.5 | 0.95 | 1.05 | 0.31 | 0.83 | 0.4 | 0.13 | 0.13 | 0.22 | 0.18 | 1.3 | NA | 0.42 | 0.49 | 1.14 | 2.95
| Analytical results for Ca in OREAS 550b (Certified Value 1.34 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.3167 | 1.31 | 1.36 | 1.33 | 1.3288 | 1.143 | 1.32 | 1.38 | 1.3396 | 1.36 | 1.34 | 1.36 | 1.36 | 1.34 | 1.3208 | 1.27 | 1.328 | 1.3634 | 1.3564 | 1.29 | 1.39 | 1.29 | 1.343 | 1
| 1 | 2 | 1.3167 | 1.32 | 1.35 | 1.33 | 1.3272 | 1.234 | 1.35 | 1.31 | 1.3178 | 1.38 | 1.32 | 1.34 | 1.36 | 1.35 | 1.3382 | 1.31 | 1.345 | 1.3772 | 1.4484 | 1.3 | 1.4 | 1.28 | 1.327 | #N/A
| 1 | 3 | 1.3367 | 1.32 | 1.39 | 1.32 | 1.3269 | 1.259 | 1.34 | 1.36 | 1.3096 | 1.36 | 1.32 | 1.28 | 1.45 | 1.35 | 1.3524 | 1.32 | 1.344 | 1.3648 | 1.4506 | 1.31 | 1.4 | 1.31 | 1.318 | 16
| 1 | 4 | 1.3458 | 1.32 | 1.36 | 1.33 | 1.3286 | 1.338 | 1.34 | 1.32 | 1.3341 | 1.37 | 1.33 | 1.26 | 1.42 | 1.35 | 1.3587 | 1.33 | 1.331 | 1.376 | 1.4215 | 1.27 | 1.41 | 1.34 | 1.338 | 1.3386
| 1 | 5 | 1.2978 | 1.32 | 1.37 | 1.33 | 1.3286 | 1.485 | 1.35 | 1.37 | 1.3399 | 1.32 | 1.3 | 1.29 | 1.42 | 1.37 | 1.3425 | 1.35 | 1.358 | 1.318 | 1.4101 | 1.3 | 1.39 | 1.31 | 1.347 | 13
| 1 | 6 | 1.3916 | 1.28 | 1.38 | 1.34 | 1.3271 | 0.926 | 1.33 | 1.34 | 1.3171 | 1.38 | 1.32 | 1.3 | 1.45 | 1.36 | 1.3736 | 1.31 | 1.347 | 1.3626 | 1.4109 | 1.32 | 1.39 | 1.26 | 1.346
| Mean |  | 1.3342 | 1.3117 | 1.3683 | 1.33 | 1.3279 | 1.2308 | 1.3383 | 1.3467 | 1.3264 | 1.3617 | 1.3217 | 1.305 | 1.41 | 1.3533 | 1.3477 | 1.315 | 1.3422 | 1.3603 | 1.4163 | 1.2983 | 1.3967 | 1.2983 | 1.3365
| Median |  | 1.3267 | 1.32 | 1.365 | 1.33 | 1.3279 | 1.2465 | 1.34 | 1.35 | 1.3259 | 1.365 | 1.32 | 1.295 | 1.42 | 1.35 | 1.3474 | 1.315 | 1.3445 | 1.3641 | 1.4162 | 1.3 | 1.395 | 1.3 | 1.3405
| Std Dev. |  | 0.032781 | 0.016021 | 0.01472 | 0.0063246 | 0.00088468 | 0.18843 | 0.01169 | 0.028048 | 0.013127 | 0.022286 | 0.013292 | 0.037815 | 0.040988 | 0.010328 | 0.01818 | 0.026646 | 0.011053 | 0.02171 | 0.034321 | 0.017224 | 0.008165 | 0.027869 | 0.01164
| Rel.Std.Dev. |  | 2.457% | 1.221% | 1.076% | 0.4755% | 0.06662% | 15.31% | 0.8735% | 2.083% | 0.9897% | 1.637% | 1.006% | 2.898% | 2.907% | 0.7631% | 1.349% | 2.026% | 0.8235% | 1.596% | 2.423% | 1.327% | 0.5846% | 2.147% | 0.871%
| PDM3 |  | -0.3258% | -2.01% | 2.223% | -0.6408% | -0.8001% | -8.049% | -0.01822% | 0.6043% | -0.9134% | 1.725% | -1.263% | -2.508% | 5.336% | 1.102% | 0.6806% | -1.761% | 0.2682% | 1.625% | 5.808% | -3.006% | 4.34% | -3.006% | -0.1552%
| Z-Score (Absolute) |  | 0.07 | 0.78 | 1 | 0.2 | 0.27 | 3.31 | 0.06 | 0.32 | 0.32 | 0.79 | 0.47 | 0.99 | 2.31 | 0.53 | 0.35 | 0.67 | 0.18 | 0.75 | 2.5 | 1.2 | 1.89 | 1.2 | 0
| Analytical results for Cd in OREAS 550b (Certified Value 0.2 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.23 | 0.31 | < 0.3 | 0.2 | 0.25 | < 0.5 | 0.22 | 0.18503 | 0.19 | 0.22 | 0.17 | 0.2 | 0.18 | < 0.5 | 0.25 | 0.19 | 0.2 | 0.2 | 0.2 | 0.14 | 0.2 |  |  | 1
| 1 | 2 | 0.2 | 0.35 | < 0.3 | 0.19 | 0.25 | < 0.5 | 0.19 | 0.22944 | 0.21 | 0.18 | 0.19 | 0.17 | 0.19 | < 0.5 | 0.25 | 0.22 | 0.22 | 0.2 | 0.2 | 0.15 | 0.21 |  |  | 8
| 1 | 3 | 0.21 | 0.4 | < 0.3 | 0.19 | 0.29 | < 0.5 | 0.21 | 0.2294 | 0.2 | 0.21 | 0.2 | 0.2 | 0.19 | < 0.5 | 0.25 | 0.2 | 0.2 | 0.2 | 0.4 | 0.16 | 0.21 |  |  | 16
| 1 | 4 | 0.22 | 0.24 | < 0.3 | 0.19 | 0.23 | < 0.5 | 0.19 | 0.20624 | 0.21 | 0.18 | 0.19 | 0.21 | 0.18 | < 0.5 | 0.23 | 0.22 | 0.23 | 0.2 | 0.2 | 0.14 | 0.2 |  |  | 0.20152
| 1 | 5 | 0.18 | 0.34 | < 0.3 | 0.19 | 0.25 | < 0.5 | 0.22 | 0.18262 | 0.19 | 0.19 | 0.2 | 0.2 | 0.19 | < 0.5 | 0.23 | 0.22 | 0.22 | 0.2 | 0.3 | 0.17 | 0.2 |  |  | 14
| 1 | 6 | 0.22 | 0.26 | < 0.3 | 0.21 | 0.23 | < 0.5 | 0.2 | 0.25109 | 0.2 | 0.2 | 0.2 | 0.22 | 0.19 | < 0.5 | 0.23 | 0.2 | 0.22 | 0.2 | 0.4 | 0.15 | 0.2
| Mean |  | 0.21 | 0.31667 |  | 0.195 | 0.25 |  | 0.205 | 0.21397 | 0.2 | 0.19667 | 0.19167 | 0.2 | 0.18667 |  | 0.24 | 0.20833 | 0.215 | 0.2 | 0.28333 | 0.15167 | 0.20333
| Median |  | 0.215 | 0.325 |  | 0.19 | 0.25 |  | 0.205 | 0.21782 | 0.2 | 0.195 | 0.195 | 0.2 | 0.19 |  | 0.24 | 0.21 | 0.22 | 0.2 | 0.25 | 0.15 | 0.2
| Std Dev. |  | 0.017889 | 0.059554 |  | 0.0083666 | 0.021909 |  | 0.013784 | 0.027334 | 0.0089443 | 0.01633 | 0.01169 | 0.016733 | 0.005164 |  | 0.010954 | 0.013292 | 0.012247 | 3.0405e-17 | 0.098319 | 0.01169 | 0.005164
| Rel.Std.Dev. |  | 8.518% | 18.81% |  | 4.291% | 8.764% |  | 6.724% | 12.77% | 4.472% | 8.303% | 6.099% | 8.367% | 2.766% |  | 4.564% | 6.38% | 5.696% | 1.52e-14% | 34.7% | 7.708% | 2.54%
| PDM3 |  | 4.209% | 57.14% |  | -3.234% | 24.06% |  | 1.728% | 6.179% | -0.7531% | -2.407% | -4.888% | -0.7531% | -7.37% |  | 19.1% | 3.382% | 6.69% | -0.7531% | 40.6% | -24.74% | 0.901%
| Z-Score (Absolute) |  | 0.34 | 7.53 | 3.71 | 0.67 | 3.03 | 3.03 | 0 | 0.6 | 0.34 | 0.56 | 0.9 | 0.34 | 1.24 | 3.03 | 2.36 | 0.22 | 0.67 | NA | NA | 3.6 | 0.11
Results from laboratories 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 550b (Certified Value 120 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
0
| 1 | 1 | 118.19 | 114.98 | 132 | 154.192 | 95.41 | 106 | 99.21 | 133 | 123 | 115.5 | 125.5 | 120.5 | 121.8108 | 123 | 120 | 124.82 | 128.5 | 121 | 113 | 121 | 86.6 |  |  | 1
| 1 | 2 | 116.32 | 114.76 | 121 | 154.116 | 99.37 | 113 | 100.79 | 137 | 120.5 | 115 | 120 | 127.5 | 123.0013 | 125 | 118 | 126.25 | 136.5 | 117 | 103 | 124 | 95.24 |  |  | 9
| 1 | 3 | 117.15 | 117.18 | 132 | 154.18 | 107 | 108 | 98.88 | 134.5 | 119 | 118 | 129 | 126.5 | 124.9215 | 128 | 112 | 123.45 | 137.9 | 117 | 106 | 122 | 96.97 |  |  | 16
| 1 | 4 | 116.85 | 114.98 | 121 | 154.124 | 99.07 | 107 | 100.99 | 133.5 | 123 | 112.5 | 126.5 | 124.5 | 125.6906 | 126 | 118 | 127.89 | 133 | 114 | 105 | 117 | 91.38 |  |  | 120.0018
| 1 | 5 | 117.18 | 113.72 | 119 | 154.184 | 107 | 110 | 98.51 | 128.5 | 121.5 | 116 | 129.5 | 124.5 | 124.242 | 126 | 118 | 122.93 | 132.4 | 117 | 114 | 120 | 95.52 |  |  | 15
| 1 | 6 | 117.12 | 113.92 | 125 | 154.112 | 102 | 108 | 101.55 | 138 | 95.1 | 113.5 | 128.5 | 125.5 | 124.9777 | 126 | 110 | 124.41 | 133.9 | 118 | 111 | 118 | 93.2
| Mean |  | 117.135 | 114.9233 | 125 | 154.1513 | 101.6417 | 108.6667 | 99.9883 | 134.0833 | 117.0167 | 115.0833 | 126.5 | 124.8333 | 124.1073 | 125.6667 | 116 | 124.9583 | 133.7 | 117.3333 | 108.6667 | 120.3333 | 93.1517
| Median |  | 117.135 | 114.87 | 123 | 154.152 | 100.685 | 108 | 100 | 134 | 121 | 115.25 | 127.5 | 125 | 124.5817 | 126 | 118 | 124.615 | 133.45 | 117 | 108.5 | 120.5 | 94.22
| Std Dev. |  | 0.61007 | 1.2313 | 5.7619 | 0.037644 | 4.651 | 2.5033 | 1.2731 | 3.3677 | 10.8454 | 1.9343 | 3.5355 | 2.4221 | 1.4452 | 1.633 | 4 | 1.8424 | 3.3088 | 2.2509 | 4.5898 | 2.582 | 3.7581
| Rel.Std.Dev. |  | 0.5208% | 1.071% | 4.61% | 0.02442% | 4.576% | 2.304% | 1.273% | 2.512% | 9.268% | 1.681% | 2.795% | 1.94% | 1.164% | 1.299% | 3.448% | 1.474% | 2.475% | 1.918% | 4.224% | 2.146% | 4.034%
| PDM3 |  | -2.389% | -4.232% | 4.165% | 28.46% | -15.3% | -9.446% | -16.68% | 11.73% | -2.488% | -4.099% | 5.415% | 4.026% | 3.421% | 4.721% | -3.335% | 4.13% | 11.42% | -2.224% | -9.446% | 0.2763% | -22.37%
| Z-Score (Absolute) |  | 0.02 | 0.21 | 0.67 | 3.24 | 1.38 | 0.76 | 1.53 | 1.47 | 0.03 | 0.2 | 0.81 | 0.66 | 0.6 | 0.73 | 0.12 | 0.67 | 1.44 | 0 | 0.76 | 0.26 | 2.13
| Analytical results for Co in OREAS 550b (Certified Value 0.143 wt.%) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.14305 | 0.12446 | 0.147 | 0.144 | 0.12558 | 0.1149 | 0.1383 | 0.1495 | 0.14232 | 0.1355 | 0.144 | 0.1425 | 0.1435 | 0.141 | 0.14113 | 0.1428 | 0.146 | 0.13252 | 0.14094 | 0.14219 | 0.14026 | 0.155 | 0.1437 | 1
| 1 | 2 | 0.1433 | 0.12645 | 0.146 | 0.145 | 0.12517 | 0.1181 | 0.1389 | 0.1474 | 0.14716 | 0.1395 | 0.1415 | 0.1395 | 0.137 | 0.144 | 0.14118 | 0.1453 | 0.147 | 0.13154 | 0.14812 | 0.14271 | 0.14204 | 0.154 | 0.142 | 29
| 1 | 3 | 0.14472 | 0.1262 | 0.147 | 0.143 | 0.12572 | 0.1239 | 0.1382 | 0.1546 | 0.14338 | 0.137 | 0.1425 | 0.1335 | 0.1445 | 0.1435 | 0.14317 | 0.1442 | 0.146 | 0.13103 | 0.15109 | 0.14176 | 0.14066 | 0.15 | 0.1392 | 16
| 1 | 4 | 0.14266 | 0.12476 | 0.144 | 0.145 | 0.12599 | 0.1154 | 0.1383 | 0.1493 | 0.14809 | 0.1365 | 0.1435 | 0.1325 | 0.141 | 0.1425 | 0.1444 | 0.1456 | 0.148 | 0.13438 | 0.14842 | 0.13871 | 0.14089 | 0.149 | 0.1433 | 0.14306
| 1 | 5 | 0.14267 | 0.12763 | 0.144 | 0.144 | 0.1263 | 0.1274 | 0.1382 | 0.1472 | 0.14778 | 0.134 | 0.1405 | 0.1345 | 0.1415 | 0.1435 | 0.14333 | 0.1458 | 0.15 | 0.13025 | 0.14674 | 0.14564 | 0.14143 | 0.151 | 0.143 | 16
| 1 | 6 | 0.14521 | 0.12135 | 0.147 | 0.145 | 0.12573 | 0.1231 | 0.138 | 0.1502 | 0.14386 | 0.1385 | 0.142 | 0.1375 | 0.1445 | 0.1425 | 0.14732 | 0.1441 | 0.146 | 0.13385 | 0.1481 | 0.14534 | 0.14205 | 0.153 | 0.1444
| Mean |  | 0.1436 | 0.12514 | 0.14583 | 0.14433 | 0.12575 | 0.12047 | 0.13832 | 0.1497 | 0.14543 | 0.13683 | 0.14233 | 0.13667 | 0.142 | 0.14283 | 0.14342 | 0.14463 | 0.14717 | 0.13226 | 0.14724 | 0.14273 | 0.14122 | 0.152 | 0.1426
| Median |  | 0.14317 | 0.12548 | 0.1465 | 0.1445 | 0.12573 | 0.1206 | 0.13825 | 0.1494 | 0.14551 | 0.13675 | 0.14225 | 0.136 | 0.1425 | 0.143 | 0.14325 | 0.14475 | 0.1465 | 0.13203 | 0.14811 | 0.14245 | 0.14116 | 0.152 | 0.14315
| Std Dev. |  | 0.0010943 | 0.0021914 | 0.001472 | 0.0008165 | 0.00038013 | 0.0050804 | 0.00030605 | 0.0026833 | 0.0025275 | 0.0019916 | 0.001291 | 0.0038687 | 0.0028636 | 0.0010801 | 0.0023006 | 0.0011466 | 0.0016021 | 0.0016224 | 0.0033961 | 0.0025563 | 0.00074155 | 0.0023664 | 0.001845
| Rel.Std.Dev. |  | 0.762% | 1.751% | 1.009% | 0.5657% | 0.3023% | 4.217% | 0.2213% | 1.792% | 1.738% | 1.456% | 0.907% | 2.831% | 2.017% | 0.7562% | 1.604% | 0.7928% | 1.089% | 1.227% | 2.307% | 1.791% | 0.5251% | 1.557% | 1.294%
| PDM3 |  | 0.3805% | -12.52% | 1.94% | 0.8919% | -12.1% | -15.79% | -3.314% | 4.643% | 1.659% | -4.351% | -0.5061% | -4.467% | -0.7391% | -0.1566% | 0.2546% | 1.102% | 2.873% | -7.546% | 2.92% | -0.2323% | -1.283% | 6.251% | -0.3197%
| Z-Score (Absolute) |  | 0.19 | 3.81 | 0.67 | 0.35 | 3.68 | 4.83 | 0.96 | 1.51 | 0.59 | 1.28 | 0.08 | 1.31 | 0.16 | 0.02 | 0.15 | 0.41 | 0.96 | 2.27 | 0.98 | 0 | 0.33 | 2.01 | 0.03
| Analytical results for Cr in OREAS 550b (Certified Value 72 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 76 | 43 | 57 | 70 | 79 | 70 | 66 | 74.1653 | 75 | 75 | 71 | 61 | 71 | 72.6134 | 64.9 | 78 | 80 | 65.7 | 65 | 71 | 62 | 85 |  | 1
| 1 | 2 | 78 | 219 | 65 | 69 | 84 | 72 | 69 | 68.6015 | 80 | 78 | 72 | 67 | 69 | 72.0853 | 65.9 | 78 | 81 | 69.6 | 68 | 72 | 64 | 80 |  | #N/A
| 1 | 3 | 79 | 47 | 68 | 69 | 89 | 74 | 65 | 66.533 | 77 | 81 | 69 | 68 | 71 | 70.7406 | 68.3 | 79 | 80 | 69.7 | 66 | 72 | 63 | 80 |  | 16
| 1 | 4 | 78 | 50 | 66 | 70 | 89 | 74 | 70 | 76.6786 | 75 | 75 | 66 | 67 | 70 | 69.1884 | 68.1 | 80 | 80 | 68.8 | 63 | 71 | 62 | 81 |  | 71.9608
| 1 | 5 | 76 | 47 | 62 | 69 | 95 | 72 | 66 | 68.9795 | 71 | 72 | 69 | 68 | 71 | 72.6146 | 71.1 | 77 | 82 | 68.2 | 64 | 72 | 63 | 80 |  | 17
| 1 | 6 | 82 | 43 | 73 | 68 | 96 | 72 | 69 | 78.9666 | 77 | 76 | 71 | 64 | 75 | 67.3956 | 69.5 | 79 | 80 | 68.2 | 65 | 72 | 59 | 59
| Mean |  | 78.1667 | 74.8333 | 65.1667 | 69.1667 | 88.6667 | 72.3333 | 67.5 | 72.3207 | 75.8333 | 76.1667 | 69.6667 | 65.8333 | 71.1667 | 70.773 | 67.9667 | 78.5 | 80.5 | 68.3667 | 65.1667 | 71.6667 | 62.1667 | 77.5
| Median |  | 78 | 47 | 65.5 | 69 | 89 | 72 | 67.5 | 71.5724 | 76 | 75.5 | 70 | 67 | 71 | 71.413 | 68.2 | 78.5 | 80 | 68.5 | 65 | 72 | 62.5 | 80
| Std Dev. |  | 2.2286 | 70.6779 | 5.4191 | 0.75277 | 6.4704 | 1.5055 | 2.0736 | 5.0011 | 2.9944 | 3.0605 | 2.1602 | 2.7869 | 2.0412 | 2.1175 | 2.2792 | 1.0488 | 0.83666 | 1.4597 | 1.7224 | 0.5164 | 1.7224 | 9.2682
| Rel.Std.Dev. |  | 2.851% | 94.45% | 8.316% | 1.088% | 7.297% | 2.081% | 3.072% | 6.915% | 3.949% | 4.018% | 3.101% | 4.233% | 2.868% | 2.992% | 3.353% | 1.336% | 1.039% | 2.135% | 2.643% | 0.7206% | 2.771% | 11.96%
| PDM3 |  | 8.624% | 3.992% | -9.441% | -3.883% | 23.22% | 0.5177% | -6.199% | 0.5002% | 5.381% | 5.845% | -3.188% | -8.515% | -1.104% | -1.651% | -5.55% | 9.087% | 11.87% | -4.995% | -9.441% | -0.4088% | -13.61% | 7.698%
| Z-Score (Absolute) |  | 1.09 | 0.55 | 1.01 | 0.36 | 2.79 | 0.15 | 0.63 | 0.15 | 0.71 | 0.77 | 0.28 | 0.9 | 0.04 | 0.1 | 0.56 | 1.15 | 1.47 | 0.49 | 1.01 | 0.04 | 1.5 | 0.98
| Analytical results for Cs in OREAS 550b (Certified Value 2.22 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.27 | 2.4 | 2.28 | 2.2 | 3.0102 | 1.94 | 2.3 | 2.34 | 2.2496 | 2.21 | 2.23 | 2.02 | 2.21 | 2.07 | 2.1454 | 2.07 | 2.23 | 2.35 | NA | 2.6 | 2.2 | 1.7 |  | 1
| 1 | 2 | 2.29 | 2.6 | 2.31 | 2.2 | 3.0531 | 2.04 | 2.54 | 2.41 | 2.3269 | 2.23 | 2.18 | 1.98 | 2.1 | 2.16 | 2.1654 | 2.11 | 2.15 | 2.28 | NA | 2.3 | 2.2 | 1.73 |  | 12
| 1 | 3 | 2.29 | 2.3 | 2.32 | 2.2 | 3.0707 | 2.14 | 2.21 | 2.25 | 2.288 | 2.15 | 2.18 | 2.07 | 2.21 | 2.14 | 2.1322 | 2.15 | 2.2 | 2.26 | NA | 2.3 | 2.2 | 1.77 |  | 16
| 1 | 4 | 2.28 | 2.3 | 2.28 | 2.2 | 3.0938 | 1.99 | 2.35 | 2.43 | 2.2657 | 2.18 | 2.22 | 1.97 | 2.18 | 2.13 | 2.2809 | 2.12 | 2.15 | 2.33 | NA | 2.2 | 2.2 | 1.71 |  | 2.2165
| 1 | 5 | 2.24 | 2.3 | 2.22 | 2.2 | 3.0687 | 2.21 | 2.34 | 2.28 | 2.2496 | 2.15 | 2.2 | 2.02 | 2.21 | 2.14 | 2.164 | 2.13 | 2.18 | 2.22 | NA | 2.4 | 2.2 | 1.72 |  | 18
| 1 | 6 | 2.26 | 2.2 | 2.36 | 2.2 | 3.0161 | 2.15 | 2.51 | 2.33 | 2.3064 | 2.24 | 2.01 | 2 | 2.22 | 2.11 | 2.2185 | 2.11 | 2.13 | 2.34 | NA | 2.7 | 2.3 | 1.69
| Mean |  | 2.2717 | 2.35 | 2.295 | 2.2 | 3.0521 | 2.0783 | 2.375 | 2.34 | 2.281 | 2.1933 | 2.17 | 2.01 | 2.1883 | 2.125 | 2.1844 | 2.115 | 2.1733 | 2.2967 |  | 2.4167 | 2.2167 | 1.72
| Median |  | 2.275 | 2.3 | 2.295 | 2.2 | 3.0609 | 2.09 | 2.345 | 2.335 | 2.2768 | 2.195 | 2.19 | 2.01 | 2.21 | 2.135 | 2.1647 | 2.115 | 2.165 | 2.305 |  | 2.35 | 2.2 | 1.715
| Std Dev. |  | 0.019408 | 0.13784 | 0.047223 | 4.8648e-16 | 0.032915 | 0.10458 | 0.12661 | 0.070427 | 0.03165 | 0.039328 | 0.080994 | 0.035777 | 0.04535 | 0.031464 | 0.0557 | 0.026646 | 0.037238 | 0.05164 |  | 0.19408 | 0.040825 | 0.028284
| Rel.Std.Dev. |  | 0.8543% | 5.866% | 2.058% | 2.211e-14% | 1.078% | 5.032% | 5.331% | 3.01% | 1.388% | 1.793% | 3.732% | 1.78% | 2.072% | 1.481% | 2.55% | 1.26% | 1.713% | 2.248% |  | 8.031% | 1.842% | 1.644%
| PDM3 |  | 2.488% | 6.022% | 3.541% | -0.745% | 37.7% | -6.234% | 7.15% | 5.571% | 2.911% | -1.046% | -2.098% | -9.317% | -1.271% | -4.129% | -1.449% | -4.58% | -1.948% | 3.616% |  | 9.03% | 0.006934% | -22.4%
| Z-Score (Absolute) |  | 0.57 | 1.19 | 0.75 | 0 | 6.76 | 0.97 | 1.39 | 1.11 | 0.64 | 0.05 | 0.24 | 1.51 | 0.09 | 0.59 | 0.12 | 0.67 | 0.21 | 0.77 | NA | 1.72 | 0.13 | 3.81
| Analytical results for Cu in OREAS 550b (Certified Value 1.84 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*MS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*AAS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.8346 | 1.811 | 1.77 | 1.81 | 1.7812 | > 1 | 1.87 | 1.838 | 1.8833 | 1.855 | 1.83 | 1.83 | 1.835 | 1.825 | 1.8354 | 1.812 | 1.83 | 1.8673 | 1.7738 | 1.829 | 1.89 | 1.899 | 1.7936 | 1
| 1 | 2 | 1.8399 | 1.796 | 1.75 | 1.84 | 1.7811 | > 1 | 1.87 | 1.844 | 1.8903 | 1.835 | 1.825 | 1.82 | 1.84 | 1.785 | 1.8636 | 1.841 | 1.84 | 1.8888 | 1.8674 | 1.85 | 1.86 | 1.866 | 1.8175 | 30
| 1 | 3 | 1.8788 | 1.802 | 1.73 | 1.84 | 1.7813 | > 1 | 1.88 | 1.856 | 1.8832 | 1.85 | 1.81 | 1.82 | 1.845 | 1.805 | 1.8746 | 1.83 | 1.82 | 1.8417 | 1.9002 | 1.833 | 1.88 | 1.821 | 1.8006 | 16
| 1 | 4 | 1.8604 | 1.789 | 1.74 | 1.83 | 1.7814 | > 1 | 1.86 | 1.844 | 1.8875 | 1.845 | 1.8 | 1.835 | 1.845 | 1.82 | 1.882 | 1.85 | 1.84 | 1.8926 | 1.8837 | 1.831 | 1.87 | 1.851 | 1.7954 | 1.8401
| 1 | 5 | 1.773 | 1.817 | 1.73 | 1.84 | 1.7813 | > 1 | 1.88 | 1.833 | 1.893 | 1.85 | 1.815 | 1.82 | 1.84 | 1.8 | 1.8587 | 1.831 | 1.86 | 1.844 | 1.8525 | 1.852 | 1.87 | 1.888 | 1.803 | 19
| 1 | 6 | 1.8665 | 1.765 | 1.73 | 1.84 | 1.7814 | > 1 | 1.87 | 1.859 | 1.8804 | 1.835 | 1.885 | 1.835 | 1.82 | 1.81 | 1.9051 | 1.839 | 1.83 | 1.8703 | 1.8503 | 1.846 | 1.88 | 1.852 | 1.7932
| Mean |  | 1.8422 | 1.7967 | 1.7417 | 1.8333 | 1.7813 |  | 1.8717 | 1.8457 | 1.8863 | 1.845 | 1.8275 | 1.8267 | 1.8375 | 1.8075 | 1.8699 | 1.8338 | 1.8367 | 1.8674 | 1.8547 | 1.8402 | 1.875 | 1.8628 | 1.8006
| Median |  | 1.8501 | 1.799 | 1.735 | 1.84 | 1.7813 |  | 1.87 | 1.844 | 1.8854 | 1.8475 | 1.82 | 1.825 | 1.84 | 1.8075 | 1.8691 | 1.835 | 1.835 | 1.8688 | 1.86 | 1.8395 | 1.875 | 1.859 | 1.798
| Std Dev. |  | 0.037711 | 0.01849 | 0.016021 | 0.012111 | 0.0001169 |  | 0.0075277 | 0.010093 | 0.0047946 | 0.0083666 | 0.030125 | 0.0075277 | 0.0093541 | 0.014405 | 0.023521 | 0.012952 | 0.013663 | 0.021514 | 0.04392 | 0.010304 | 0.010488 | 0.028138 | 0.0091886
| Rel.Std.Dev. |  | 2.047% | 1.029% | 0.9199% | 0.6606% | 0.006563% |  | 0.4022% | 0.5468% | 0.2542% | 0.4535% | 1.648% | 0.4121% | 0.5091% | 0.7969% | 1.258% | 0.7063% | 0.7439% | 1.152% | 2.368% | 0.5599% | 0.5594% | 1.511% | 0.5103%
| PDM3 |  | 0.1129% | -2.361% | -5.35% | -0.3683% | -3.197% |  | 1.715% | 0.3019% | 2.51% | 0.2657% | -0.6853% | -0.7306% | -0.1419% | -1.772% | 1.62% | -0.3411% | -0.1872% | 1.485% | 0.7906% | 0.003042% | 1.896% | 1.235% | -2.15%
| Z-Score (Absolute) |  | 0.11 | 1.43 | 3.29 | 0.19 | 1.95 | NA | 1.11 | 0.23 | 1.61 | 0.21 | 0.38 | 0.41 | 0.05 | 1.06 | 1.05 | 0.17 | 0.07 | 0.97 | 0.54 | 0.05 | 1.22 | 0.81 | 1.3
| Analytical results for Dy in OREAS 550b (Certified Value 3.81 ppm) | [24 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 3.86 | 3.5 | 4.4 | 4.7601 | 3.36 | 3.9293 | 3.72 | 3.75 | [15 empty] | 1
| 1 | 2 | 3.86 | 3.8 | 4.3 | 4.8064 | 3.7 | 3.9881 | 3.81 | 3.65 | [15 empty] | 13
| 1 | 3 | 3.87 | 3.4 | 4.4 | 4.8242 | 3.81 | 3.8635 | 3.9 | 3.62 | [15 empty] | 16
| 1 | 4 | 3.78 | 3.5 | 4.4 | 4.7374 | 3.54 | 4.1666 | 3.71 | 3.57 | [15 empty] | 3.8059
| 1 | 5 | 3.85 | 3.4 | 4.1 | 4.813 | 3.73 | 3.876 | 3.74 | 3.71 | [15 empty] | 20
| 1 | 6 | 3.86 | 3.1 | 4.3 | 4.8005 | 3.64 | 4.0027 | 3.65 | 3.73
| Mean |  | 3.8467 | 3.45 | 4.3167 | 4.7903 | 3.63 | 3.971 | 3.755 | 3.6717
| Median |  | 3.86 | 3.45 | 4.35 | 4.8034 | 3.67 | 3.9587 | 3.73 | 3.68
| Std Dev. |  | 0.033267 | 0.22583 | 0.1169 | 0.033889 | 0.16025 | 0.11122 | 0.087807 | 0.069976
| Rel.Std.Dev. |  | 0.8648% | 6.546% | 2.708% | 0.7075% | 4.415% | 2.801% | 2.338% | 1.906%
| PDM3 |  | 1.072% | -9.35% | 13.42% | 25.87% | -4.621% | 4.34% | -1.336% | -3.526%
| Z-Score (Absolute) |  | 0.18 | 1.39 | 2.04 | 3.91 | 0.68 | 0.67 | 0.18 | 0.51
| Analytical results for Er in OREAS 550b (Certified Value 2.18 ppm) | [24 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 2.15 | 2.1 | 2.4 | 2.8965 | 1.75 | 2.326 | 2.16 | 2.07 | [15 empty] | 1
| 1 | 2 | 2.11 | 2.2 | 2.4 | 2.9465 | 1.81 | 2.3714 | 2.22 | 2.1 | [15 empty] | 14
| 1 | 3 | 2.19 | 2 | 2.5 | 2.9542 | 2.04 | 2.3114 | 2.26 | 2.08 | [15 empty] | 16
| 1 | 4 | 2.11 | 2.1 | 2.5 | 2.9334 | 1.88 | 2.3953 | 2.19 | 2.09 | [15 empty] | 2.1789
| 1 | 5 | 2.13 | 2 | 2.5 | 2.9182 | 2.08 | 2.3487 | 2.28 | 2.03 | [15 empty] | 21
| 1 | 6 | 2.22 | 1.9 | 2.5 | 2.8977 | 2.07 | 2.3391 | 2.18 | 2.12
| Mean |  | 2.1517 | 2.05 | 2.4667 | 2.9244 | 1.9383 | 2.3486 | 2.215 | 2.0817
| Median |  | 2.14 | 2.05 | 2.5 | 2.9258 | 1.96 | 2.3439 | 2.205 | 2.085
| Std Dev. |  | 0.044907 | 0.10488 | 0.05164 | 0.024439 | 0.14359 | 0.030606 | 0.047223 | 0.030605
| Rel.Std.Dev. |  | 2.087% | 5.116% | 2.094% | 0.8357% | 7.408% | 1.303% | 2.132% | 1.47%
| PDM3 |  | -1.248% | -5.914% | 13.21% | 34.22% | -11.04% | 7.792% | 1.659% | -4.46%
| Z-Score (Absolute) |  | 0.14 | 0.6 | 1.28 | 3.35 | 1.11 | 0.75 | 0.14 | 0.46
| Analytical results for Eu in OREAS 550b (Certified Value 0.82 ppm) | [24 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.88 | 0.7 | 0.9 | 1.3756 | 0.8 | 0.84307 | 0.79 | 0.75 | [15 empty] | 1
| 1 | 2 | 0.81 | 0.8 | 0.92 | 1.3685 | 0.77 | 0.88983 | 0.83 | 0.79 | [15 empty] | 15
| 1 | 3 | 0.82 | 0.8 | 0.91 | 1.3769 | 0.84 | 0.83012 | 0.82 | 0.81 | [15 empty] | 16
| 1 | 4 | 0.84 | 0.8 | 0.89 | 1.3191 | 0.81 | 0.90017 | 0.81 | 0.78 | [15 empty] | 0.82332
| 1 | 5 | 0.82 | 0.7 | 0.86 | 1.3399 | 0.84 | 0.85523 | 0.84 | 0.8 | [15 empty] | 22
| 1 | 6 | 0.82 | 0.7 | 0.92 | 1.3562 | 0.81 | 0.85082 | 0.82 | 0.81
| Mean |  | 0.83167 | 0.75 | 0.9 | 1.356 | 0.81167 | 0.86154 | 0.81833 | 0.79
| Median |  | 0.82 | 0.75 | 0.905 | 1.3623 | 0.81 | 0.85303 | 0.82 | 0.795
| Std Dev. |  | 0.025626 | 0.054772 | 0.022804 | 0.02278 | 0.026394 | 0.02748 | 0.017224 | 0.022804
| Rel.Std.Dev. |  | 3.081% | 7.303% | 2.534% | 1.68% | 3.252% | 3.19% | 2.105% | 2.887%
| PDM3 |  | 1.014% | -8.905% | 9.314% | 64.7% | -1.415% | 4.643% | -0.6051% | -4.046%
| Z-Score (Absolute) |  | 0.13 | 1.41 | 1.41 | 10.01 | 0.25 | 0.69 | 0.13 | 0.66
| Analytical results for Fe in OREAS 550b (Certified Value 2.74 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.78 | 2.69 | 2.83 | 2.85 | 2.7143 | 2.7 | 2.65 | 2.7 | 2.834 | 2.67 | 2.73 | 2.71 | 2.67 | 2.67 | 2.7357 | 2.58 | 2.78 | 2.81 | 2.6094 | 2.67 | 2.74 | 2.34 | 2.76 | 1
| 1 | 2 | 2.78 | 2.79 | 2.85 | 2.85 | 2.7072 | 2.86 | 2.77 | 2.67 | 2.7886 | 2.74 | 2.68 | 2.69 | 2.68 | 2.7 | 2.7041 | 2.59 | 2.8 | 2.83 | 2.7769 | 2.66 | 2.76 | 2.37 | 2.73 | #N/A
| 1 | 3 | 2.83 | 2.75 | 2.87 | 2.86 | 2.7142 | 2.6 | 2.66 | 2.71 | 2.7502 | 2.68 | 2.7 | 2.55 | 2.85 | 2.7 | 2.7381 | 2.63 | 2.8 | 2.78 | 2.7888 | 2.68 | 2.75 | 2.38 | 2.73 | 16
| 1 | 4 | 2.82 | 2.76 | 2.76 | 2.84 | 2.7092 | 2.68 | 2.77 | 2.66 | 2.7904 | 2.67 | 2.71 | 2.54 | 2.81 | 2.67 | 2.7682 | 2.6 | 2.81 | 2.85 | 2.7457 | 2.59 | 2.8 | 2.51 | 2.76 | 2.7398
| 1 | 5 | 2.72 | 2.79 | 2.82 | 2.84 | 2.7114 | 2.47 | 2.8 | 2.73 | 2.8322 | 2.61 | 2.68 | 2.58 | 2.81 | 2.71 | 2.7372 | 2.66 | 2.84 | 2.74 | 2.7225 | 2.67 | 2.78 | 2.49 | 2.76 | 23
| 1 | 6 | 2.92 | 2.65 | 2.95 | 2.87 | 2.7119 | 2.99 | 2.8 | 2.65 | 2.7397 | 2.72 | 2.64 | 2.64 | 2.85 | 2.7 | 2.7759 | 2.6 | 2.81 | 2.84 | 2.7128 | 2.7 | 2.76 | 2.35 | 2.79
| Mean |  | 2.8083 | 2.7383 | 2.8467 | 2.8517 | 2.7114 | 2.7167 | 2.7417 | 2.6867 | 2.7892 | 2.6817 | 2.69 | 2.6183 | 2.7783 | 2.6917 | 2.7432 | 2.61 | 2.8067 | 2.8083 | 2.726 | 2.6617 | 2.765 | 2.4067 | 2.755
| Median |  | 2.8 | 2.755 | 2.84 | 2.85 | 2.7117 | 2.69 | 2.77 | 2.685 | 2.7895 | 2.675 | 2.69 | 2.61 | 2.81 | 2.7 | 2.7376 | 2.6 | 2.805 | 2.82 | 2.7341 | 2.67 | 2.76 | 2.375 | 2.76
| Std Dev. |  | 0.067057 | 0.056716 | 0.062823 | 0.01169 | 0.002796 | 0.18511 | 0.068532 | 0.031411 | 0.039586 | 0.04535 | 0.030984 | 0.072503 | 0.082077 | 0.017224 | 0.025857 | 0.029665 | 0.019664 | 0.041673 | 0.06435 | 0.037639 | 0.021679 | 0.073937 | 0.022583
| Rel.Std.Dev. |  | 2.388% | 2.071% | 2.207% | 0.41% | 0.1031% | 6.814% | 2.5% | 1.169% | 1.419% | 1.691% | 1.152% | 2.769% | 2.954% | 0.6399% | 0.9426% | 1.137% | 0.7006% | 1.484% | 2.361% | 1.414% | 0.7841% | 3.072% | 0.8197%
| PDM3 |  | 2.501% | -0.05377% | 3.9% | 4.083% | -1.038% | -0.8446% | 0.06789% | -1.94% | 1.802% | -2.122% | -1.818% | -4.434% | 1.406% | -1.757% | 0.1229% | -4.738% | 2.44% | 2.501% | -0.5033% | -2.852% | 0.9195% | -12.16% | 0.5545%
| Z-Score (Absolute) |  | 0.93 | 0 | 1.44 | 1.5 | 0.36 | 0.29 | 0.04 | 0.69 | 0.67 | 0.75 | 0.64 | 1.59 | 0.53 | 0.62 | 0.06 | 1.7 | 0.91 | 0.93 | 0.16 | 1.02 | 0.35 | 4.4 | 0.22
| Analytical results for Ga in OREAS 550b (Certified Value 22.7 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 22.31 | 22.68 | 23.3 | 22 | 13.14 | 20.6 | 22.1 | 22.3668 | 22.5 | 23.8 | 22.9 | 21.2 | 21 | 23.4206 | 23.7 | 24 | 23.69 | 20.32 | 25 | 22 | 20.15 |  |  | 1
| 1 | 2 | 22.02 | 23.1 | 23 | 22.1 | 13.57 | 23 | 22.6 | 23.8872 | 23.2 | 23.9 | 22.5 | 20.4 | 21.6 | 23.1961 | 23.85 | 23.69 | 23.23 | 21.88 | 24 | 21.7 | 20.21 |  |  | 17
| 1 | 3 | 22.24 | 22.68 | 23.8 | 22 | 14.17 | 21.9 | 22.3 | 22.2406 | 23 | 23.3 | 23 | 21.9 | 21.7 | 23.722 | 24.49 | 22.8 | 23.11 | 18.33 | 24 | 21.3 | 20 |  |  | 16
| 1 | 4 | 22.27 | 22.68 | 23.7 | 22.1 | 13.07 | 23.4 | 22.6 | 23.5503 | 23 | 23.9 | 22.7 | 21.9 | 21.9 | 23.6775 | 23.89 | 23.71 | 23.47 | 16.14 | 25 | 21.4 | 20.37 |  |  | 22.7422
| 1 | 5 | 22.3 | 23.2 | 23.4 | 22.3 | 14.3 | 23.5 | 22.5 | 22.7566 | 22.8 | 23.7 | 23.5 | 21.6 | 21.7 | 23.559 | 24.5 | 23.38 | 22.94 | 19.87 | 25 | 21.5 | 20.3 |  |  | 24
| 1 | 6 | 22.25 | 22.36 | 23.9 | 22 | 14.09 | 22.5 | 22.2 | 23.2009 | 23.5 | 23.5 | 22.9 | 21.4 | 21.6 | 23.7006 | 24.31 | 22.38 | 23.71 | 20.3 | 26 | 22 | 20.16
| Mean |  | 22.2317 | 22.7833 | 23.5167 | 22.0833 | 13.7233 | 22.4833 | 22.3833 | 23.0004 | 23 | 23.6833 | 22.9167 | 21.4 | 21.5833 | 23.546 | 24.1233 | 23.3267 | 23.3583 | 19.4733 | 24.8333 | 21.65 | 20.1983
| Median |  | 22.26 | 22.68 | 23.55 | 22.05 | 13.83 | 22.75 | 22.4 | 22.9787 | 23 | 23.75 | 22.9 | 21.5 | 21.65 | 23.6183 | 24.1 | 23.535 | 23.35 | 20.085 | 25 | 21.6 | 20.185
| Std Dev. |  | 0.10722 | 0.31149 | 0.34303 | 0.1169 | 0.53984 | 1.098 | 0.2137 | 0.65837 | 0.34059 | 0.24014 | 0.33714 | 0.56214 | 0.30605 | 0.2052 | 0.352 | 0.61786 | 0.316 | 1.9884 | 0.75277 | 0.30166 | 0.12891
| Rel.Std.Dev. |  | 0.4823% | 1.367% | 1.459% | 0.5294% | 3.934% | 4.884% | 0.9547% | 2.862% | 1.481% | 1.014% | 1.471% | 2.627% | 1.418% | 0.8715% | 1.459% | 2.649% | 1.353% | 10.21% | 3.031% | 1.393% | 0.6382%
| PDM3 |  | -2.245% | 0.181% | 3.406% | -2.897% | -39.66% | -1.138% | -1.578% | 1.135% | 1.134% | 4.138% | 0.7673% | -5.902% | -5.096% | 3.534% | 6.073% | 2.57% | 2.709% | -14.37% | 9.195% | -4.802% | -11.19%
| Z-Score (Absolute) |  | 0.51 | 0 | 0.67 | 0.64 | 8.33 | 0.28 | 0.37 | 0.2 | 0.2 | 0.83 | 0.12 | 1.27 | 1.1 | 0.7 | 1.23 | 0.5 | 0.53 | 3.04 | 1.88 | 1.04 | 2.38
| Analytical results for Gd in OREAS 550b (Certified Value 4.67 ppm) | [24 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 4.5 | 4.9 | 5.4 | 7.3107 | 4.05 | 4.9375 | 4.51 | 3.9 | [15 empty] | 1
| 1 | 2 | 4.45 | 5.2 | 5.2 | 7.249 | 4.23 | 4.9534 | 4.92 | 4.1 | [15 empty] | 18
| 1 | 3 | 4.34 | 4.8 | 5.2 | 7.2945 | 4.66 | 4.8461 | 4.82 | 4.1 | [15 empty] | 16
| 1 | 4 | 4.38 | 4.9 | 5.3 | 7.2265 | 4.19 | 5.2161 | 4.64 | 4 | [15 empty] | 4.6715
| 1 | 5 | 4.34 | 4.7 | 5.1 | 7.2453 | 4.66 | 5.0065 | 4.79 | 4 | [15 empty] | 25
| 1 | 6 | 4.47 | 4.5 | 5.6 | 7.2578 | 4.48 | 4.9928 | 4.82 | 4.1
| Mean |  | 4.4133 | 4.8333 | 5.3 | 7.264 | 4.3783 | 4.9921 | 4.75 | 4.0333
| Median |  | 4.415 | 4.85 | 5.25 | 7.2534 | 4.355 | 4.9731 | 4.805 | 4.05
| Std Dev. |  | 0.069186 | 0.23381 | 0.17889 | 0.032044 | 0.25857 | 0.12343 | 0.14832 | 0.08165
| Rel.Std.Dev. |  | 1.568% | 4.837% | 3.375% | 0.4411% | 5.906% | 2.473% | 3.123% | 2.024%
| PDM3 |  | -5.526% | 3.465% | 13.45% | 55.5% | -6.275% | 6.862% | 1.681% | -13.66%
| Z-Score (Absolute) |  | 0.64 | 0.07 | 0.87 | 4.21 | 0.7 | 0.34 | 0.07 | 1.29
| Analytical results for Ge in OREAS 550b (Indicative Value 0.23 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 18 | 19 | 21 | 25 | 26 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 1.7 | 0.19 | < 0.1 | 1.5 | 0.48 | 0.3 | 0.2468 | 0.15 | 0.19 | 0.12 | 0.16 | 0.24 | 0.27815 | 1.1 | < 0.02 | 3.8 | 1.71 | < 0.2 | [5 empty] | 1
| 1 | 2 | 1.8 | 0.25 | < 0.1 | 1.5 | 0.47 | 0.4 | 0.29119 | 0.14 | 0.22 | 0.14 | 0.16 | 0.24 | 0.26273 | 1.3 | 0.06 | 3.8 | 1.66 | < 0.2 | [5 empty] | 2
| 1 | 3 | 1.7 | 0.17 | < 0.1 | 1.5 | 0.42 | 0.4 | 0.24529 | 0.14 | 0.17 | 0.15 | 0.19 | 0.22 | 0.28979 | 1.3 | 0.07 | 3.7 | 1.65 | < 0.2 | [5 empty] | 16
| 1 | 4 | 1.8 | 0.2 | < 0.1 | 1.4 | 0.46 | 0.3 | 0.28914 | 0.16 | 0.18 | 0.14 | 0.19 | 0.26 | 0.3371 | 1.4 | 0.07 | 3.5 | 1.66 | < 0.2 | [5 empty] | 0.23053
| 1 | 5 | 1.7 | 0.22 | < 0.1 | 1.5 | 0.44 | 0.4 | 0.26361 | 0.15 | 0.19 | 0.16 | 0.18 | 0.23 | 0.32619 | 1.3 | 0.06 | 3.6 | 1.74 | < 0.2 | [5 empty] | 8
| 1 | 6 | 1.8 | 0.19 | < 0.1 | 1.5 | 0.46 | 0.3 | 0.30602 | 0.15 | 0.17 | 0.13 | 0.17 | 0.23 | 0.31299 | 1.4 | 0.07 | 3.6 | 1.74 | < 0.2
| Mean |  | 1.75 | 0.20333 |  | 1.4833 | 0.455 | 0.35 | 0.27368 | 0.14833 | 0.18667 | 0.14 | 0.175 | 0.23667 | 0.30116 | 1.3 | 0.066 | 3.6667 | 1.6933
| Median |  | 1.75 | 0.195 |  | 1.5 | 0.46 | 0.35 | 0.27638 | 0.15 | 0.185 | 0.14 | 0.175 | 0.235 | 0.30139 | 1.3 | 0.07 | 3.65 | 1.685
| Std Dev. |  | 0.054772 | 0.028048 |  | 0.040825 | 0.021679 | 0.054772 | 0.025383 | 0.0075277 | 0.018619 | 0.014142 | 0.013784 | 0.013663 | 0.028961 | 0.10954 | 0.0054772 | 0.12111 | 0.041793
| Rel.Std.Dev. |  | 3.13% | 13.79% |  | 2.752% | 4.765% | 15.65% | 9.275% | 5.075% | 9.974% | 10.1% | 7.877% | 5.773% | 9.617% | 8.427% | 8.299% | 3.303% | 2.468%
| PDM3 |  | 659.1% | -11.8% |  | 543.4% | 97.37% | 51.82% | 18.72% | -35.66% | -19.03% | -39.27% | -24.09% | 2.662% | 30.64% | 463.9% | -71.37% | 1491% | 634.5%
| Z-Score (Absolute) |  | 7.46 | 0.26 | 1.02 | 6.13 | 1 | 0.47 | 0.09 | 0.53 | 0.34 | 0.57 | 0.4 | 0.09 | 0.23 | 5.21 | 0.99 | 17.02 | 7.17 | 0.77
| Analytical results for Hf in OREAS 550b (Certified Value 3.83 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 3.78 | 3.07 | 4 | 4 | 4.87 | 3.6 | 3.98 | 3.9844 | 3.8 | 4.1 | 3.8 | 3.5 | 3.6 | 3.9421 | 4.01 | 3.68 | 3.81 | 2.34 | 3.5 | 3.9 | 3.99 | 3.77 |  | 1
| 1 | 2 | 3.71 | 3.41 | 4.1 | 3.9 | 3.99 | 4.1 | 4.25 | 3.9458 | 3.9 | 4 | 3.7 | 3.8 | 3.7 | 3.7303 | 3.89 | 3.73 | 3.8 | 2.92 | 3.5 | 3.7 | 4.08 | 3.76 |  | 22
| 1 | 3 | 3.73 | 3.17 | 4.2 | 3.8 | 3.91 | 3.6 | 4.09 | 3.9286 | 3.8 | 4 | 3.7 | 4.1 | 3.7 | 3.7923 | 3.77 | 3.89 | 3.71 | 3.01 | 3.5 | 3.7 | 3.93 | 3.89 |  | 16
| 1 | 4 | 3.72 | 3.1 | 4.1 | 4 | 3.36 | 3.8 | 4.22 | 3.8733 | 3.8 | 4.2 | 3.7 | 3.9 | 3.7 | 3.8221 | 3.88 | 3.69 | 3.78 | 4.24 | 3.4 | 3.7 | 3.78 | 3.82 |  | 3.8276
| 1 | 5 | 3.62 | 3.14 | 4 | 3.9 | 3.61 | 3.7 | 3.97 | 4.0183 | 3.9 | 4 | 3.7 | 4 | 3.7 | 3.7999 | 3.88 | 3.66 | 3.66 | 3.32 | 3.4 | 3.8 | 4.02 | 3.87 |  | 26
| 1 | 6 | 3.84 | 2.89 | 4.1 | 3.9 | 3.59 | 3.5 | 4.19 | 3.9167 | 4 | 4 | 3.6 | 4.1 | 3.7 | 3.875 | 3.71 | 3.76 | 3.69 | 3.33 | 3.6 | 3.7 | 3.9 | 3.72
| Mean |  | 3.7333 | 3.13 | 4.0833 | 3.9167 | 3.8883 | 3.7167 | 4.1167 | 3.9445 | 3.8667 | 4.05 | 3.7 | 3.9 | 3.6833 | 3.827 | 3.8567 | 3.735 | 3.7417 | 3.1933 | 3.4833 | 3.75 | 3.95 | 3.805
| Median |  | 3.725 | 3.12 | 4.1 | 3.9 | 3.76 | 3.65 | 4.14 | 3.9372 | 3.85 | 4 | 3.7 | 3.95 | 3.7 | 3.811 | 3.88 | 3.71 | 3.745 | 3.165 | 3.5 | 3.7 | 3.96 | 3.795
| Std Dev. |  | 0.073666 | 0.16864 | 0.075277 | 0.075277 | 0.53293 | 0.2137 | 0.12226 | 0.051271 | 0.08165 | 0.083666 | 0.063246 | 0.22804 | 0.040825 | 0.073282 | 0.10463 | 0.084083 | 0.063061 | 0.62717 | 0.075277 | 0.083666 | 0.10507 | 0.066558
| Rel.Std.Dev. |  | 1.973% | 5.388% | 1.844% | 1.922% | 13.71% | 5.75% | 2.97% | 1.3% | 2.112% | 2.066% | 1.709% | 5.847% | 1.108% | 1.915% | 2.713% | 2.251% | 1.685% | 19.64% | 2.161% | 2.231% | 2.66% | 1.749%
| PDM3 |  | -2.463% | -18.23% | 6.682% | 2.327% | 1.587% | -2.898% | 7.552% | 3.055% | 1.021% | 5.811% | -3.333% | 1.892% | -3.769% | -0.01623% | 0.7596% | -2.419% | -2.245% | -16.57% | -8.994% | -2.027% | 3.198% | -0.5902%
| Z-Score (Absolute) |  | 0.56 | 4.63 | 1.8 | 0.68 | 0.49 | 0.67 | 2.03 | 0.87 | 0.34 | 1.58 | 0.78 | 0.57 | 0.89 | 0.07 | 0.27 | 0.55 | 0.5 | 4.2 | 2.24 | 0.44 | 0.9 | 0.07
| Analytical results for Hg in OREAS 550b (Indicative Value 0.056 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [20 empty] | 1
| No | No | 15 | 19 | 25 | [20 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [20 empty] | 3
3
| 1 | 1 | < 1 | < 0.01 | 0.07 | [20 empty] | 1
| 1 | 2 | < 1 | 0.12 | 0.08 | [20 empty] | 1
| 1 | 3 | < 1 | 0.05 | 0.08 | [20 empty] | 16
| 1 | 4 | < 1 | 0.02 | 0.06 | [20 empty] | 0.05625
| 1 | 5 | < 1 | 0.03 | 0.06 | [20 empty] | 7
| 1 | 6 | < 1 | 0.05 | 0.05
| Mean |  |  | 0.054 | 0.066667
| Median |  |  | 0.05 | 0.065
| Std Dev. |  |  | 0.039115 | 0.012111
| Rel.Std.Dev. |  |  | 72.44% | 18.17%
| PDM3 |  |  | -4% | 18.52%
| Z-Score (Absolute) |  | 14.03 | 0.67 | 0
| Analytical results for Ho in OREAS 550b (Certified Value 0.73 ppm) | [24 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.73 | 0.6 | 0.9 | 0.93504 | 0.67 | 0.76633 | 0.75 | 0.49 | [15 empty] | 1
| 1 | 2 | 0.72 | 0.7 | 0.8 | 0.95232 | 0.7 | 0.7818 | 0.77 | 0.5 | [15 empty] | 5
| 1 | 3 | 0.69 | 0.6 | 0.9 | 0.95368 | 0.75 | 0.76904 | 0.77 | 0.51 | [15 empty] | 16
| 1 | 4 | 0.72 | 0.6 | 0.8 | 0.94528 | 0.69 | 0.81889 | 0.73 | 0.49 | [15 empty] | 0.73479
| 1 | 5 | 0.72 | 0.6 | 0.8 | 0.92992 | 0.76 | 0.78433 | 0.75 | 0.51 | [15 empty] | 27
| 1 | 6 | 0.73 | 0.6 | 0.9 | 0.93872 | 0.73 | 0.79466 | 0.72 | 0.5
| Mean |  | 0.71833 | 0.61667 | 0.85 | 0.94249 | 0.71667 | 0.78584 | 0.74833 | 0.5
| Median |  | 0.72 | 0.6 | 0.85 | 0.942 | 0.715 | 0.78307 | 0.75 | 0.5
| Std Dev. |  | 0.01472 | 0.040825 | 0.054772 | 0.009564 | 0.03559 | 0.019241 | 0.020412 | 0.0089443
| Rel.Std.Dev. |  | 2.049% | 6.62% | 6.444% | 1.015% | 4.966% | 2.449% | 2.728% | 1.789%
| PDM3 |  | -2.24% | -16.08% | 15.68% | 28.27% | -2.467% | 6.947% | 1.843% | -31.95%
| Z-Score (Absolute) |  | 0.12 | 0.93 | 0.93 | 1.67 | 0.13 | 0.42 | 0.12 | 1.86
| Analytical results for In in OREAS 550b (Certified Value 0.96 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 0.99 | 1.11 | 0.9 | 0.95 | 0.89 | < 2 | 0.88 | 0.94295 | 1.01 | 0.971 | 0.905 | 0.992 | 0.954 | 0.9192 | 0.94 | 0.985 | 0.98 | 0.218 | 0.9 | 0.95 | 1.05 | 0.811 |  | 1
| 1 | 2 | 0.97 | 1.24 | 0.9 | 0.96 | 0.9 | < 2 | 0.91 | 0.98012 | 1.005 | 0.942 | 0.894 | 0.963 | 0.965 | 0.90488 | 0.93 | 0.942 | 0.98 | 0.209 | 0.91 | 0.95 | 1.01 | 0.797 |  | 6
| 1 | 3 | 1 | 1.11 | 0.9 | 0.95 | 0.948 | < 2 | 0.89 | 0.93035 | 1 | 0.952 | 0.949 | 1.05 | 0.954 | 0.92168 | 0.96 | 0.967 | 0.99 | 0.207 | 0.9 | 0.94 | 1.05 | 0.785 |  | 16
| 1 | 4 | 1 | 1.12 | 0.9 | 0.96 | 0.853 | < 2 | 0.91 | 0.98653 | 0.998 | 0.956 | 0.903 | 1.04 | 0.965 | 0.96866 | 0.97 | 0.963 | 0.99 | 0.068 | 0.9 | 0.93 | 0.945 | 0.807 |  | 0.96135
| 1 | 5 | 0.95 | 1.09 | 0.9 | 0.96 | 0.943 | < 2 | 0.88 | 1.0103 | 1.03 | 0.961 | 0.919 | 1.01 | 0.98 | 0.95324 | 0.98 | 0.984 | 0.95 | 0.226 | 0.93 | 0.98 | 1 | 0.796 |  | 28
| 1 | 6 | 0.97 | 1.04 | 1 | 0.96 | 0.874 | < 2 | 0.92 | 0.99042 | 1.045 | 0.938 | 0.905 | 1.025 | 0.959 | 0.9659 | 0.95 | 0.945 | 0.96 | 0.237 | 0.93 | 0.99 | 0.95 | 0.806
| Mean |  | 0.98 | 1.1183 | 0.91667 | 0.95667 | 0.90133 |  | 0.89833 | 0.97344 | 1.0147 | 0.95333 | 0.9125 | 1.0133 | 0.96283 | 0.93893 | 0.955 | 0.96433 | 0.975 | 0.19417 | 0.91167 | 0.95667 | 1.0008 | 0.80033
| Median |  | 0.98 | 1.11 | 0.9 | 0.96 | 0.895 |  | 0.9 | 0.98332 | 1.0075 | 0.954 | 0.905 | 1.0175 | 0.962 | 0.93746 | 0.955 | 0.965 | 0.98 | 0.2135 | 0.905 | 0.95 | 1.005 | 0.8015
| Std Dev. |  | 0.02 | 0.066156 | 0.040825 | 0.005164 | 0.037755 |  | 0.017224 | 0.030491 | 0.018779 | 0.012193 | 0.019593 | 0.032222 | 0.0097451 | 0.027064 | 0.018708 | 0.018414 | 0.016432 | 0.062799 | 0.01472 | 0.023381 | 0.046089 | 0.0095429
| Rel.Std.Dev. |  | 2.041% | 5.916% | 4.454% | 0.5398% | 4.189% |  | 1.917% | 3.132% | 1.851% | 1.279% | 2.147% | 3.18% | 1.012% | 2.882% | 1.959% | 1.909% | 1.685% | 32.34% | 1.615% | 2.444% | 4.605% | 1.192%
| PDM3 |  | 1.939% | 16.33% | -4.648% | -0.4876% | -6.243% |  | -6.555% | 1.257% | 5.546% | -0.8344% | -5.082% | 5.407% | 0.1538% | -2.333% | -0.661% | 0.3098% | 1.419% | -79.8% | -5.169% | -0.4876% | 4.107% | -16.75%
| Z-Score (Absolute) |  | 0.38 | 2.62 | 0.65 | 0 | 0.9 | 0.7 | 0.94 | 0.27 | 0.94 | 0.05 | 0.71 | 0.92 | 0.1 | 0.29 | 0.03 | 0.12 | 0.3 | 12.34 | 0.73 | 0 | 0.71 | 2.53
| Analytical results for K in OREAS 550b (Certified Value 2.02 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.0007 | 1.45 | 2.25 | 2 | 1.8394 | 1.63 | 1.99 | 2.02 | 2.0376 | 2.01 | 1.97 | 1.95 | 1.97 | 2 | 1.9514 | 2 | 2.03 | 2.0252 | 1.9018 | 2 | 1.96 | 2 | 2.1 | 1
| 1 | 2 | 2.0253 | 1.66 | 2.14 | 2.01 | 1.8427 | 1.79 | 2.01 | 1.95 | 2.0211 | 2.04 | 1.95 | 1.95 | 1.98 | 2.01 | 2.0043 | 2.01 | 2.03 | 2.0431 | 2.0398 | 2.09 | 1.97 | 1.98 | 2.1 | #N/A
| 1 | 3 | 2.0485 | 1.43 | 2.2 | 1.99 | 1.8388 | 1.86 | 1.99 | 2.03 | 1.9944 | 2 | 1.96 | 1.85 | 2.13 | 2.03 | 2.0041 | 2.03 | 2.04 | 2.019 | 2.0364 | 2.11 | 1.97 | 2.01 | 2.07 | 16
| 1 | 4 | 2.0473 | 1.62 | 2.29 | 2.01 | 1.8389 | 1.94 | 2.06 | 1.97 | 2.0195 | 2.01 | 1.96 | 1.84 | 2.08 | 2 | 2.0122 | 2.01 | 2.05 | 2.0391 | 1.9979 | 1.93 | 1.98 | 2.04 | 2.12 | 2.0174
| 1 | 5 | 1.979 | 1.67 | 2.17 | 2.01 | 1.8469 | 2.17 | 2.02 | 2.03 | 2.0499 | 1.96 | 1.94 | 1.87 | 2.11 | 2.03 | 1.9947 | 2.04 | 2.06 | 1.9883 | 1.9703 | 2.07 | 1.97 | 1.98 | 2.12 | 29
| 1 | 6 | 2.116 | 1.49 | 2.22 | 2.02 | 1.8373 | 1.25 | 2.02 | 1.99 | 1.9902 | 2.04 | 1.92 | 1.91 | 2.13 | 2.01 | 2.041 | 2 | 2.04 | 2.0462 | 1.9835 | 2.11 | 1.97 | 1.99 | 2.11
| Mean |  | 2.0361 | 1.5533 | 2.2117 | 2.0067 | 1.8407 | 1.7733 | 2.015 | 1.9983 | 2.0188 | 2.01 | 1.95 | 1.895 | 2.0667 | 2.0133 | 2.0013 | 2.015 | 2.0417 | 2.0268 | 1.9883 | 2.0517 | 1.97 | 2 | 2.1033
| Median |  | 2.0363 | 1.555 | 2.21 | 2.01 | 1.8392 | 1.825 | 2.015 | 2.005 | 2.0203 | 2.01 | 1.955 | 1.89 | 2.095 | 2.01 | 2.0042 | 2.01 | 2.04 | 2.0321 | 1.9907 | 2.08 | 1.97 | 1.995 | 2.105
| Std Dev. |  | 0.047526 | 0.10893 | 0.054191 | 0.010328 | 0.0035365 | 0.3122 | 0.025884 | 0.033714 | 0.023397 | 0.029665 | 0.017889 | 0.048888 | 0.073394 | 0.013663 | 0.029113 | 0.016432 | 0.01169 | 0.021629 | 0.050769 | 0.072226 | 0.0063246 | 0.022804 | 0.018619
| Rel.Std.Dev. |  | 2.334% | 7.013% | 2.45% | 0.5147% | 0.1921% | 17.61% | 1.285% | 1.687% | 1.159% | 1.476% | 0.9174% | 2.58% | 3.551% | 0.6786% | 1.455% | 0.8155% | 0.5726% | 1.067% | 2.553% | 3.52% | 0.321% | 1.14% | 0.8852%
| PDM3 |  | 0.9293% | -23% | 9.63% | -0.5313% | -8.76% | -12.1% | -0.1183% | -0.9444% | 0.06836% | -0.3661% | -3.34% | -6.067% | 2.443% | -0.2009% | -0.7983% | -0.1183% | 1.204% | 0.4675% | -1.443% | 1.699% | -2.349% | -0.8618% | 4.26%
| Z-Score (Absolute) |  | 0.67 | 11.78 | 5.2 | 0.09 | 4.37 | 6.11 | 0.13 | 0.3 | 0.23 | 0 | 1.55 | 2.97 | 1.46 | 0.09 | 0.22 | 0.13 | 0.82 | 0.43 | 0.56 | 1.08 | 1.03 | 0.26 | 2.41
| Analytical results for La in OREAS 550b (Certified Value 61 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
1
| 1 | 1 | 60.67 | 58.9 | 64.1 | 77.6136 | 53.26 | 56 | 49 | 68.7 | 68.7 | 60.8 | 63.1 | 61.9 | 59.9997 | 61.9 | 62.45 | 65.19 | 49.1 | 61.8 | 69 | 61.6 | 50.2 |  |  | 1
| 1 | 2 | 59.73 | 58.9 | 62.4 | 77.5072 | 55.6 | 58 | 48.6 | 70.2 | 64.6 | 60.7 | 60.6 | 63.9 | 60.3624 | 62.8 | 62.15 | 66.01 | 51.5 | 60.1 | 69 | 63.4 | 55.4 |  |  | 7
| 1 | 3 | 60.23 | 59.6 | 64.5 | 77.6352 | 59.63 | 56 | 48.2 | 69.3 | 64.7 | 60.6 | 66.9 | 65.1 | 60.3489 | 64.7 | 57.91 | 64.94 | 44.6 | 60.1 | 70 | 63.2 | 54.7 |  |  | 16
| 1 | 4 | 62.01 | 59.1 | 63.4 | 77.5632 | 55.98 | 56 | 49 | 68.6 | 65.9 | 59.9 | 67.4 | 65.1 | 60.7777 | 64.3 | 61.27 | 66.36 | 38.9 | 58.4 | 70 | 60.5 | 53.1 |  |  | 60.6183
| 1 | 5 | 58.12 | 59 | 61.9 | 77.6168 | 60.07 | 60 | 48.3 | 65.9 | 65.2 | 61.6 | 66.5 | 64.3 | 59.5867 | 64.2 | 60.59 | 64.61 | 48.6 | 60 | 70 | 62.3 | 54.5 |  |  | 30
| 1 | 6 | 61.81 | 58.7 | 65.1 | 77.5584 | 57.7 | 56 | 48.9 | 71.1 | 52 | 60.2 | 68.6 | 65.6 | 60.7288 | 63.5 | 57.87 | 65.21 | 49.3 | 62.1 | 69 | 60.1 | 54.1
| Mean |  | 60.4283 | 59.0333 | 63.5667 | 77.5824 | 57.04 | 57 | 48.6667 | 68.9667 | 63.5167 | 60.6333 | 65.5167 | 64.3167 | 60.3007 | 63.5667 | 60.3733 | 65.3867 | 47 | 60.4167 | 69.5 | 61.85 | 53.6667
| Median |  | 60.45 | 58.95 | 63.75 | 77.5884 | 56.84 | 56 | 48.75 | 69 | 64.95 | 60.65 | 66.7 | 64.7 | 60.3556 | 63.85 | 60.93 | 65.2 | 48.85 | 60.1 | 69.5 | 61.95 | 54.3
| Std Dev. |  | 1.437 | 0.30768 | 1.2388 | 0.048051 | 2.6004 | 1.6733 | 0.3559 | 1.7773 | 5.841 | 0.58195 | 3.0314 | 1.3333 | 0.45106 | 1.0577 | 2.0325 | 0.66461 | 4.5581 | 1.3556 | 0.54772 | 1.3693 | 1.8597
| Rel.Std.Dev. |  | 2.378% | 0.5212% | 1.949% | 0.06193% | 4.559% | 2.936% | 0.7313% | 2.577% | 9.196% | 0.9598% | 4.627% | 2.073% | 0.748% | 1.664% | 3.367% | 1.016% | 9.698% | 2.244% | 0.7881% | 2.214% | 3.465%
| PDM3 |  | -0.3134% | -2.615% | 4.864% | 27.99% | -5.903% | -5.969% | -19.72% | 13.77% | 4.781% | 0.02483% | 8.081% | 6.101% | -0.5239% | 4.864% | -0.4041% | 7.866% | -22.47% | -0.3326% | 14.65% | 2.032% | -11.47%
| Z-Score (Absolute) |  | 0.04 | 0.3 | 0.55 | 3.18 | 0.67 | 0.68 | 2.25 | 1.56 | 0.54 | 0 | 0.92 | 0.69 | 0.06 | 0.55 | 0.05 | 0.89 | 2.56 | 0.04 | 1.66 | 0.23 | 1.31
| Analytical results for Li in OREAS 550b (Certified Value 77 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 73.7 | 117 | 74 | 74.1 | 78.67 | 62 | 76 | 74 | 78.2936 | 79.4 | 75.7 | 80 | 78.3 | 79.7 | 74.7767 | 75.7 | 75.9 | 75.9 | 83 | 78.3 | 74.4 | 77.6 | 82 | 1
| 1 | 2 | 75.2 | 118.9 | 74 | 74.1 | 79.43 | 68 | 78 | 76 | 81.3048 | 80.7 | 75.1 | 79.7 | 75.3 | 80 | 78.2685 | 76.5 | 75.2 | 77.2 | 79 | 80.1 | 76.2 | 75.3 | 80 | #N/A
| 1 | 3 | 75.3 | 118.2 | 76 | 74.1 | 81.39 | 68 | 77 | 73 | 76.7173 | 79.1 | 75.5 | 75.6 | 80.6 | 80.7 | 76.0344 | 77 | 75.9 | 75.7 | 75 | 79.1 | 79.6 | 72.5 | 77 | 16
| 1 | 4 | 75.9 | 119 | 75 | 74.1 | 81.1 | 71 | 78 | 77 | 83.2916 | 79.3 | 75.9 | 75.2 | 79 | 79.4 | 76.9992 | 75.7 | 75.9 | 77.4 | 75 | 79.7 | 78.8 | 74.1 | 81 | 76.9276
| 1 | 5 | 74.4 | 121.1 | 73 | 75.1 | 80.53 | 78 | 79 | 73 | 79.3873 | 77.7 | 75.2 | 76 | 79.9 | 80.6 | 75.2309 | 77.6 | 76.7 | 76.2 | 73 | 77.7 | 77.5 | 75.3 | 80 | 31
| 1 | 6 | 76.2 | 115.6 | 75 | 75.1 | 79.28 | 46 | 78 | 75 | 81.5304 | 80.3 | 75.4 | 78 | 80.7 | 80.4 | 78.1099 | 76.6 | 76 | 77.7 | 73 | 81.8 | 73.3 | 74.8 | 81
| Mean |  | 75.1167 | 118.3 | 74.5 | 74.4333 | 80.0667 | 65.5 | 77.6667 | 74.6667 | 80.0875 | 79.4167 | 75.4667 | 77.4167 | 78.9667 | 80.1333 | 76.5699 | 76.5167 | 75.9333 | 76.6833 | 76.3333 | 79.45 | 76.6333 | 74.9333 | 80.1667
| Median |  | 75.25 | 118.55 | 74.5 | 74.1 | 79.98 | 68 | 78 | 74.5 | 80.3461 | 79.35 | 75.45 | 77 | 79.45 | 80.2 | 76.5168 | 76.55 | 75.9 | 76.7 | 75 | 79.4 | 76.85 | 75.05 | 80.5
| Std Dev. |  | 0.93256 | 1.8804 | 1.0488 | 0.5164 | 1.0962 | 10.8766 | 1.0328 | 1.633 | 2.404 | 1.0477 | 0.30111 | 2.1189 | 2.0216 | 0.52026 | 1.4654 | 0.7414 | 0.4761 | 0.85186 | 3.9328 | 1.4502 | 2.4712 | 1.6741 | 1.7224
| Rel.Std.Dev. |  | 1.241% | 1.59% | 1.408% | 0.6938% | 1.369% | 16.61% | 1.33% | 2.187% | 3.002% | 1.319% | 0.399% | 2.737% | 2.56% | 0.6492% | 1.914% | 0.9689% | 0.627% | 1.111% | 5.152% | 1.825% | 3.225% | 2.234% | 2.149%
| PDM3 |  | -2.354% | 53.78% | -3.156% | -3.242% | 4.081% | -14.86% | 0.9607% | -2.939% | 4.108% | 3.236% | -1.899% | 0.6357% | 2.651% | 4.167% | -0.4649% | -0.5342% | -1.292% | -0.3175% | -0.7725% | 3.279% | -0.3825% | -2.592% | 4.211%
| Z-Score (Absolute) |  | 0.52 | 14.29 | 0.73 | 0.75 | 1.18 | 3.82 | 0.35 | 0.67 | 1.18 | 0.95 | 0.4 | 0.27 | 0.8 | 1.2 | 0.02 | 0.04 | 0.24 | 0.02 | 0.1 | 0.97 | 0 | 0.58 | 1.21
| Analytical results for Lu in OREAS 550b (Certified Value 0.33 ppm) | [24 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 15 | 17 | 25 | 26 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.33 | 0.3 | 0.4 | 0.4278 | 0.32 | 0.31 | 0.36366 | 0.33 | 0.34 | 0.24 | [13 empty] | 1
| 1 | 2 | 0.32 | 0.3 | 0.3 | 0.4216 | 0.32 | 0.32 | 0.35723 | 0.33 | 0.32 | 0.25 | [13 empty] | 8
| 1 | 3 | 0.33 | 0.3 | 0.3 | 0.442 | 0.36 | 0.3 | 0.32604 | 0.35 | 0.33 | 0.26 | [13 empty] | 16
| 1 | 4 | 0.33 | 0.3 | 0.3 | 0.41792 | 0.33 | 0.32 | 0.3563 | 0.34 | 0.34 | 0.26 | [13 empty] | 0.33275
| 1 | 5 | 0.31 | 0.3 | 0.3 | 0.43016 | 0.37 | 0.3 | 0.36053 | 0.35 | 0.33 | 0.26 | [13 empty] | 32
| 1 | 6 | 0.33 | 0.3 | 0.4 | 0.44496 | 0.34 | 0.32 | 0.34519 | 0.33 | 0.32 | 0.26
| Mean |  | 0.325 | 0.3 | 0.33333 | 0.43074 | 0.34 | 0.31167 | 0.35149 | 0.33833 | 0.33 | 0.255
| Median |  | 0.33 | 0.3 | 0.3 | 0.42898 | 0.335 | 0.315 | 0.35677 | 0.335 | 0.33 | 0.26
| Std Dev. |  | 0.0083666 | 0 | 0.05164 | 0.010825 | 0.020976 | 0.0098319 | 0.013951 | 0.0098319 | 0.0089443 | 0.0083666
| Rel.Std.Dev. |  | 2.574% | 0% | 15.49% | 2.513% | 6.169% | 3.155% | 3.969% | 2.906% | 2.71% | 3.281%
| PDM3 |  | -2.329% | -9.842% | 0.1757% | 29.45% | 2.179% | -6.336% | 5.633% | 1.678% | -0.826% | -23.37%
| Z-Score (Absolute) |  | 0.47 | NA | NA | 4.92 | 0.3 | 1.15 | 0.88 | 0.21 | 0.21 | 4.03
Results from laboratories 2 and 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 550b (Certified Value 4.31 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.2299 | 4.34 | 4.65 | 4.17 | 3.4269 | 4.094 | 4.23 | 4.16 | 4.4845 | 4.34 | 4.3 | 4.34 | 4.37 | 4.31 | 4.3044 | 4.3 | 4.328 | 4.249 | 4.2265 | 4.29 | 4.07 | 4.07 | 4.428 | 1
| 1 | 2 | 4.2996 | 4.39 | 4.55 | 4.2 | 3.4155 | 4.101 | 4.32 | 4.2 | 4.4345 | 4.41 | 4.24 | 4.33 | 4.21 | 4.34 | 4.3576 | 4.47 | 4.322 | 4.2669 | 4.4843 | 4.27 | 4.11 | 4.11 | 4.354 | #N/A
| 1 | 3 | 4.3364 | 4.43 | 4.76 | 4.18 | 3.4202 | 4.592 | 4.36 | 4.13 | 4.3641 | 4.31 | 4.26 | 4.08 | 4.51 | 4.37 | 4.3551 | 4.49 | 4.324 | 4.2056 | 4.5168 | 4.29 | 4.13 | 4.2 | 4.323 | 16
| 1 | 4 | 4.3036 | 4.4 | 4.58 | 4.25 | 3.4234 | 4.436 | 4.4 | 4.23 | 4.415 | 4.33 | 4.29 | 4.06 | 4.41 | 4.34 | 4.4157 | 4.54 | 4.348 | 4.2663 | 4.4215 | 4.19 | 4.15 | 4.18 | 4.412 | 4.3121
| 1 | 5 | 4.1767 | 4.45 | 4.57 | 4.25 | 3.4259 | 4.952 | 4.44 | 4.12 | 4.4788 | 4.22 | 4.25 | 4.14 | 4.47 | 4.38 | 4.3786 | 4.6 | 4.392 | 4.1448 | 4.3905 | 4.33 | 4.13 | 4.25 | 4.415 | 33
| 1 | 6 | 4.4766 | 4.27 | 4.69 | 4.27 | 3.4118 | 4.702 | 4.36 | 4.22 | 4.3376 | 4.4 | 4.18 | 4.22 | 4.52 | 4.37 | 4.4357 | 4.48 | 4.364 | 4.2676 | 4.3996 | 4.34 | 4.09 | 4.01 | 4.427
| Mean |  | 4.3038 | 4.38 | 4.6333 | 4.22 | 3.4206 | 4.4795 | 4.3517 | 4.1767 | 4.4191 | 4.335 | 4.2533 | 4.195 | 4.415 | 4.3517 | 4.3745 | 4.48 | 4.3463 | 4.2334 | 4.4065 | 4.285 | 4.1133 | 4.1367 | 4.3932
| Median |  | 4.3016 | 4.395 | 4.615 | 4.225 | 3.4218 | 4.514 | 4.36 | 4.18 | 4.4247 | 4.335 | 4.255 | 4.18 | 4.44 | 4.355 | 4.3681 | 4.485 | 4.338 | 4.2576 | 4.4106 | 4.29 | 4.12 | 4.145 | 4.4135
| Std Dev. |  | 0.10252 | 0.065727 | 0.08165 | 0.041952 | 0.0059871 | 0.34035 | 0.072226 | 0.046762 | 0.059614 | 0.06892 | 0.04274 | 0.12194 | 0.11589 | 0.026394 | 0.046954 | 0.1006 | 0.027667 | 0.049478 | 0.10114 | 0.053572 | 0.029439 | 0.089368 | 0.043925
| Rel.Std.Dev. |  | 2.382% | 1.501% | 1.762% | 0.9941% | 0.175% | 7.598% | 1.66% | 1.12% | 1.349% | 1.59% | 1.005% | 2.907% | 2.625% | 0.6065% | 1.073% | 2.245% | 0.6366% | 1.169% | 2.295% | 1.25% | 0.7157% | 2.16% | 0.9998%
| PDM3 |  | -0.1936% | 1.573% | 7.448% | -2.137% | -20.67% | 3.881% | 0.9164% | -3.142% | 2.48% | 0.5299% | -1.364% | -2.717% | 2.385% | 0.9164% | 1.446% | 3.893% | 0.7928% | -1.827% | 2.189% | -0.6296% | -4.611% | -4.069% | 1.879%
| Z-Score (Absolute) |  | 0.39 | 0.31 | 2.66 | 1.17 | 8.58 | 1.23 | 0.05 | 1.57 | 0.67 | 0.11 | 0.86 | 1.4 | 0.64 | 0.05 | 0.26 | 1.24 | 0 | 1.05 | 0.56 | 0.57 | 2.16 | 1.94 | 0.43
| Analytical results for Mn in OREAS 550b (Certified Value 0.019 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0185 | 0.0195 | 0.0192 | 0.0195 | 0.019448 | 0.0172 | 0.0174 | 0.0183 | 0.019524 | 0.0194 | 0.0189 | 0.0191 | 0.019 | 0.0186 | 0.019181 | 0.0189 | 0.0192 | 0.02 | 0.0168 | 0.0193 | 0.0189 | 0.014 | 0.0191 | 1
| 1 | 2 | 0.0185 | 0.0205 | 0.0171 | 0.0189 | 0.018968 | 0.0187 | 0.0192 | 0.0189 | 0.019154 | 0.0192 | 0.0187 | 0.0189 | 0.0181 | 0.0189 | 0.019422 | 0.019 | 0.0195 | 0.0203 | 0.0177 | 0.0193 | 0.0192 | 0.0138 | 0.0192 | #N/A
| 1 | 3 | 0.0189 | 0.0197 | 0.0205 | 0.0189 | 0.01905 | 0.019 | 0.019 | 0.0183 | 0.01899 | 0.0187 | 0.0187 | 0.0179 | 0.0194 | 0.0188 | 0.019462 | 0.0194 | 0.0194 | 0.0201 | 0.0176 | 0.0188 | 0.0191 | 0.0143 | 0.0188 | 16
| 1 | 4 | 0.0187 | 0.0194 | 0.0175 | 0.0192 | 0.018996 | 0.02 | 0.0189 | 0.0193 | 0.019272 | 0.0189 | 0.0187 | 0.0178 | 0.019 | 0.0188 | 0.019406 | 0.0191 | 0.0195 | 0.0206 | 0.0175 | 0.0187 | 0.0192 | 0.0139 | 0.019 | 0.019058
| 1 | 5 | 0.0181 | 0.0198 | 0.0208 | 0.0193 | 0.019456 | 0.0224 | 0.0198 | 0.0182 | 0.019405 | 0.0185 | 0.0185 | 0.0181 | 0.0192 | 0.019 | 0.019416 | 0.0195 | 0.0197 | 0.0197 | 0.0174 | 0.0194 | 0.0194 | 0.0137 | 0.0192 | 34
| 1 | 6 | 0.0195 | 0.0192 | 0.0195 | 0.0196 | 0.019283 | 0.0133 | 0.0207 | 0.0192 | 0.018985 | 0.0191 | 0.0183 | 0.0185 | 0.0195 | 0.0189 | 0.019787 | 0.0195 | 0.0194 | 0.0204 | 0.0173 | 0.0197 | 0.0191 | 0.0135 | 0.0193
| Mean |  | 0.0187 | 0.019683 | 0.0191 | 0.019233 | 0.0192 | 0.018433 | 0.019167 | 0.0187 | 0.019222 | 0.018967 | 0.018633 | 0.018383 | 0.019033 | 0.018833 | 0.019446 | 0.019233 | 0.01945 | 0.020183 | 0.017383 | 0.0192 | 0.01915 | 0.013867 | 0.0191
| Median |  | 0.0186 | 0.0196 | 0.01935 | 0.01925 | 0.019167 | 0.01885 | 0.0191 | 0.0186 | 0.019213 | 0.019 | 0.0187 | 0.0183 | 0.0191 | 0.01885 | 0.019419 | 0.01925 | 0.01945 | 0.0202 | 0.01745 | 0.0193 | 0.01915 | 0.01385 | 0.01915
| Std Dev. |  | 0.00047329 | 0.0004535 | 0.0015218 | 0.00029439 | 0.00022444 | 0.0030481 | 0.001093 | 0.00049396 | 0.00021974 | 0.00033267 | 0.00020656 | 0.00053821 | 0.00050067 | 0.00013663 | 0.00019488 | 0.00026583 | 0.00016432 | 0.00031885 | 0.00031885 | 0.00037947 | 0.00016432 | 0.00027325 | 0.00017889
| Rel.Std.Dev. |  | 2.531% | 2.304% | 7.968% | 1.531% | 1.169% | 16.54% | 5.703% | 2.642% | 1.143% | 1.754% | 1.109% | 2.928% | 2.63% | 0.7254% | 1.002% | 1.382% | 0.8448% | 1.58% | 1.834% | 1.976% | 0.8581% | 1.971% | 0.9366%
| PDM3 |  | -1.878% | 3.282% | 0.2207% | 0.9203% | 0.7463% | -3.277% | 0.5705% | -1.878% | 0.8585% | -0.479% | -2.228% | -3.54% | -0.1291% | -1.179% | 2.034% | 0.9203% | 2.057% | 5.905% | -8.787% | 0.7454% | 0.483% | -27.24% | 0.2207%
| Z-Score (Absolute) |  | 1.01 | 1.48 | 0 | 0.34 | 0.25 | 1.69 | 0.17 | 1.01 | 0.31 | 0.34 | 1.18 | 1.81 | 0.17 | 0.67 | 0.87 | 0.34 | 0.89 | 2.74 | 4.34 | 0.25 | 0.13 | 13.23 | 0
| Analytical results for Mo in OREAS 550b (Certified Value 6.97 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 7.2 | 7.01 | 8 | 6.8 | 7 | 7.32 | 6.0463 | 7.02 | 6.94 | 7.03 | 7.14 | 6.38 | 6.9216 | 6.37 | 7.55 | 7.2 | 7 | 6.5 | 7.4 | 6.5 | 6.8 |  |  | 1
| 1 | 2 | 7 | 7.82 | 8 | 6.8 | 7 | 7.13 | 6.3658 | 7.26 | 6.81 | 6.89 | 6.83 | 6.71 | 6.8979 | 6.51 | 7.28 | 7 | 7.6 | 6.4 | 7.3 | 6.2 | 6.7 |  |  | 9
| 1 | 3 | 7.2 | 7.08 | 7 | 6.8 | 7 | 7.18 | 5.9296 | 7.25 | 6.73 | 6.95 | 7.36 | 6.66 | 7.0557 | 6.51 | 7.45 | 7.2 | 7.2 | 6.7 | 7.2 | 6.3 | 6.8 |  |  | 16
| 1 | 4 | 7.3 | 6.96 | 7 | 6.9 | 7 | 7.29 | 6.2682 | 7.17 | 6.78 | 6.8 | 7.14 | 6.56 | 6.9679 | 6.52 | 7.38 | 7.4 | 7.5 | 6.4 | 7.2 | 6.1 | 6.9 |  |  | 6.9731
| 1 | 5 | 7 | 7.03 | 7 | 6.9 | 7 | 7.12 | 6.4356 | 7.27 | 6.83 | 7.07 | 7.04 | 6.51 | 7.045 | 6.79 | 7.39 | 7.1 | 7.3 | 6.5 | 7.2 | 6.3 | 7 |  |  | 35
| 1 | 6 | 7.3 | 6.76 | 7 | 6.8 | 7 | 7.32 | 5.8918 | 7.26 | 6.83 | 6.81 | 7.06 | 6.49 | 7.3183 | 6.53 | 7.24 | 7.2 | 7.3 | 6.6 | 7.4 | 6.1 | 6.8
| Mean |  | 7.1667 | 7.11 | 7.3333 | 6.8333 | 7 | 7.2267 | 6.1562 | 7.205 | 6.82 | 6.925 | 7.095 | 6.5517 | 7.0344 | 6.5383 | 7.3817 | 7.1833 | 7.3167 | 6.5167 | 7.2833 | 6.25 | 6.8333
| Median |  | 7.2 | 7.02 | 7 | 6.8 | 7 | 7.235 | 6.1572 | 7.255 | 6.82 | 6.92 | 7.1 | 6.535 | 7.0065 | 6.515 | 7.385 | 7.2 | 7.3 | 6.5 | 7.25 | 6.25 | 6.8
| Std Dev. |  | 0.13663 | 0.36508 | 0.5164 | 0.05164 | 0 | 0.094163 | 0.23148 | 0.097724 | 0.069857 | 0.11203 | 0.17248 | 0.1199 | 0.15292 | 0.13688 | 0.11268 | 0.13292 | 0.2137 | 0.1169 | 0.098319 | 0.15166 | 0.10328
| Rel.Std.Dev. |  | 1.906% | 5.135% | 7.042% | 0.7557% | 0% | 1.303% | 3.76% | 1.356% | 1.024% | 1.618% | 2.431% | 1.83% | 2.174% | 2.094% | 1.526% | 1.85% | 2.921% | 1.794% | 1.35% | 2.427% | 1.511%
| PDM3 |  | 2.776% | 1.963% | 5.166% | -2.005% | 0.3855% | 3.636% | -11.71% | 3.325% | -2.196% | -0.6901% | 1.748% | -6.044% | 0.8787% | -6.235% | 5.859% | 3.015% | 4.927% | -6.546% | 4.449% | -10.37% | -2.005%
| Z-Score (Absolute) |  | 0.44 | 0.25 | 1 | 0.67 | 0.12 | 0.64 | 2.95 | 0.57 | 0.72 | 0.37 | 0.2 | 1.62 | 0 | 1.66 | 1.16 | 0.5 | 0.95 | 1.74 | 0.83 | 2.63 | 0.67
| Analytical results for Na in OREAS 550b (Certified Value 0.053 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0503 | 0.093 | 0.06 | 0.042 | 0.115 | 0.05 | 0.35 | 0.058709 | 0.05 | 0.05 | 0.06 | 0.05 | 0.07 | 0.051979 | 0.064 | 0.051 | 0.0523 | 0.0493 | 0.051 | 0.06 | 0.0426 | 0.059 |  | 1
| 1 | 2 | 0.0505 | 0.083 | 0.05 | 0.039 | 0.101 | 0.05 | 0.33 | 0.057295 | 0.05 | 0.05 | 0.06 | 0.07 | 0.07 | 0.051751 | 0.061 | 0.052 | 0.0531 | 0.0507 | 0.051 | 0.06 | 0.0415 | 0.059 |  | 21
| 1 | 3 | 0.0506 | 0.09 | 0.05 | 0.038 | 0.101 | 0.05 | 0.32 | 0.057115 | 0.05 | 0.05 | 0.05 | 0.07 | 0.07 | 0.052505 | 0.063 | 0.052 | 0.0514 | 0.0507 | 0.051 | 0.06 | 0.0413 | 0.06 |  | 16
| 1 | 4 | 0.0503 | 0.068 | 0.05 | 0.04 | 0.099 | 0.05 | 0.35 | 0.057515 | 0.05 | 0.05 | 0.05 | 0.07 | 0.07 | 0.054083 | 0.063 | 0.052 | 0.0517 | 0.0492 | 0.05 | 0.06 | 0.0402 | 0.059 |  | 0.052798
| 1 | 5 | 0.0491 | 0.083 | 0.05 | 0.039 | 0.103 | 0.05 | 0.32 | 0.058775 | 0.05 | 0.05 | 0.05 | 0.07 | 0.07 | 0.051895 | 0.063 | 0.053 | 0.0504 | 0.0491 | 0.05 | 0.06 | 0.0399 | 0.06 |  | 36
| 1 | 6 | 0.0516 | 0.083 | 0.06 | 0.039 | 0.102 | 0.05 | 0.33 | 0.056439 | 0.05 | 0.05 | 0.06 | 0.07 | 0.07 | 0.0538 | 0.064 | 0.053 | 0.0515 | 0.0488 | 0.051 | 0.05 | 0.0392 | 0.059
| Mean |  | 0.0504 | 0.083333 | 0.053333 | 0.0395 | 0.1035 | 0.05 | 0.33333 | 0.057641 | 0.05 | 0.05 | 0.055 | 0.066667 | 0.07 | 0.052669 | 0.063 | 0.052167 | 0.051733 | 0.049633 | 0.050667 | 0.058333 | 0.040783 | 0.059333
| Median |  | 0.0504 | 0.083 | 0.05 | 0.039 | 0.1015 | 0.05 | 0.33 | 0.057405 | 0.05 | 0.05 | 0.055 | 0.07 | 0.07 | 0.052242 | 0.063 | 0.052 | 0.0516 | 0.04925 | 0.051 | 0.06 | 0.04075 | 0.059
| Std Dev. |  | 0.0008 | 0.008641 | 0.005164 | 0.0013784 | 0.0057879 | 7.6012e-18 | 0.013663 | 0.00092546 | 7.6012e-18 | 7.6012e-18 | 0.0054772 | 0.008165 | 0 | 0.0010222 | 0.0010954 | 0.00075277 | 0.00090921 | 0.00084301 | 0.0005164 | 0.0040825 | 0.0012416 | 0.0005164
| Rel.Std.Dev. |  | 1.587% | 10.37% | 9.682% | 3.49% | 5.592% | 1.52e-14% | 4.099% | 1.606% | 1.52e-14% | 1.52e-14% | 9.959% | 12.25% | 0% | 1.941% | 1.739% | 1.443% | 1.757% | 1.698% | 1.019% | 6.999% | 3.044% | 0.8703%
| PDM3 |  | -4.541% | 57.83% | 1.014% | -25.19% | 96.03% | -5.299% | 531.3% | 9.174% | -5.299% | -5.299% | 4.171% | 26.27% | 32.58% | -0.2443% | 19.32% | -1.195% | -2.016% | -5.994% | -4.036% | 10.48% | -22.76% | 12.38%
| Z-Score (Absolute) |  | 0.44 | 5.11 | 0.06 | 2.27 | 8.5 | 0.51 | 47.21 | 0.78 | 0.51 | 0.51 | 0.34 | 2.3 | 2.86 | 0.06 | 1.68 | 0.14 | 0.21 | 0.57 | 0.39 | 0.9 | 2.06 | 1.07
| Analytical results for Nb in OREAS 550b (Certified Value 10.7 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
1
| 1 | 1 | 11.43 | 14.3 | 21.5 | 8.3 | 5.7 | 4.9 | 5.6 | 8.6658 | 13.4 | 13.8 | 13.8 | 9.3 | 14.2 | 13.7062 | 14.5 | 9.7 | 11.58 | NA | 10.4 | 9.8 | 11.8 | 11.9 |  | 1
| 1 | 2 | 11.13 | 15.68 | 20.9 | 6.1 | 5.6 | 4.8 | 5.3 | 8.7843 | 13.7 | 13.6 | 13.7 | 10.5 | 14.5 | 13.4708 | 15 | 9.3 | 11.08 | NA | 9.3 | 10.1 | 11.4 | 12.2 |  | 25
| 1 | 3 | 11.27 | 14.6 | 21.6 | 5.8 | 5.5 | 5.1 | 5.4 | 8.5619 | 13.9 | 13.8 | 13.9 | 11.7 | 14.5 | 13.3897 | 14.6 | 9.6 | 10.86 | NA | 10.2 | 10 | 11.2 | 11.8 |  | 16
| 1 | 4 | 11.38 | 14.41 | 21.1 | 5.6 | 5.3 | 5 | 5.2 | 8.8331 | 13.4 | 14 | 13.7 | 11.6 | 15.1 | 13.4258 | 14.7 | 9.2 | 10.59 | NA | 8.8 | 10.1 | 11.4 | 11.5 |  | 10.7128
| 1 | 5 | 11.74 | 14.07 | 21.2 | 5.2 | 5.2 | 5.2 | 5.7 | 8.4568 | 13.1 | 13.7 | 14.1 | 10.4 | 14.7 | 13.5246 | 14.5 | 9.8 | 10.75 | NA | 8.7 | 9.8 | 11.6 | 11 |  | 37
| 1 | 6 | 11.53 | 13.82 | 20.8 | 6 | 5.1 | 5.6 | 5.8 | 8.5538 | 13.6 | 13.4 | 13.4 | 10.8 | 15.2 | 13.4396 | 14.7 | 9.4 | 11.2 | NA | 9.2 | 10.2 | 11.4 | 11
| Mean |  | 11.4133 | 14.48 | 21.1833 | 6.1667 | 5.4 | 5.1 | 5.5 | 8.6426 | 13.5167 | 13.7167 | 13.7667 | 10.7167 | 14.7 | 13.4928 | 14.6667 | 9.5 | 11.01 |  | 9.4333 | 10 | 11.4667 | 11.5667
| Median |  | 11.405 | 14.355 | 21.15 | 5.9 | 5.4 | 5.05 | 5.5 | 8.6138 | 13.5 | 13.75 | 13.75 | 10.65 | 14.6 | 13.4552 | 14.65 | 9.5 | 10.97 |  | 9.25 | 10.05 | 11.4 | 11.65
| Std Dev. |  | 0.21097 | 0.64736 | 0.31885 | 1.093 | 0.23664 | 0.28284 | 0.23664 | 0.14548 | 0.27869 | 0.20412 | 0.23381 | 0.88412 | 0.38471 | 0.11401 | 0.18619 | 0.23664 | 0.35564 |  | 0.71181 | 0.16733 | 0.20656 | 0.49261
| Rel.Std.Dev. |  | 1.848% | 4.471% | 1.505% | 17.72% | 4.382% | 5.546% | 4.303% | 1.683% | 2.062% | 1.488% | 1.698% | 8.25% | 2.617% | 0.845% | 1.269% | 2.491% | 3.23% |  | 7.546% | 1.673% | 1.801% | 4.259%
| PDM3 |  | 6.54% | 35.17% | 97.74% | -42.44% | -49.59% | -52.39% | -48.66% | -19.32% | 26.17% | 28.04% | 28.51% | 0.03638% | 37.22% | 25.95% | 36.91% | -11.32% | 2.775% |  | -11.94% | -6.653% | 7.037% | 7.971%
| Z-Score (Absolute) |  | 0 | 0.9 | 2.86 | 1.54 | 1.76 | 1.85 | 1.73 | 0.81 | 0.62 | 0.67 | 0.69 | 0.2 | 0.96 | 0.61 | 0.95 | 0.56 | 0.12 | NA | 0.58 | 0.41 | 0.02 | 0.04
| Analytical results for Nd in OREAS 550b (Certified Value 48.4 ppm) | [24 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 1
2
| 1 | 1 | 48.7 | 43.1 | 52.6 | 60.2136 | 44 | 50.2946 | 49 | 35.9 | [15 empty] | 1
| 1 | 2 | 48.09 | 45.4 | 51.9 | 60.5928 | 46.6 | 50.8059 | 50.3 | 37.4 | [15 empty] | 12
| 1 | 3 | 48.08 | 42.6 | 52.3 | 60.1336 | 49.8 | 49.2163 | 50.4 | 37.6 | [15 empty] | 16
| 1 | 4 | 47.63 | 42.7 | 51.7 | 60.3688 | 45.6 | 51.8379 | 48.3 | 36.9 | [15 empty] | 48.4473
| 1 | 5 | 46.75 | 42.5 | 50.8 | 60.4712 | 50.4 | 49.9523 | 49.8 | 37.7 | [15 empty] | 38
| 1 | 6 | 49.42 | 38.4 | 52.1 | 60.8136 | 47.8 | 51.1646 | 49.2 | 37.3
| Mean |  | 48.1117 | 42.45 | 51.9 | 60.4323 | 47.3667 | 50.5453 | 49.5 | 37.1333
| Median |  | 48.085 | 42.65 | 52 | 60.42 | 47.2 | 50.5502 | 49.5 | 37.35
| Std Dev. |  | 0.91055 | 2.2634 | 0.6229 | 0.2505 | 2.4639 | 0.92768 | 0.81486 | 0.66533
| Rel.Std.Dev. |  | 1.893% | 5.332% | 1.2% | 0.4145% | 5.202% | 1.835% | 1.646% | 1.792%
| PDM3 |  | -0.6927% | -12.38% | 7.127% | 24.74% | -2.23% | 4.33% | 2.173% | -23.35%
| Z-Score (Absolute) |  | 0.19 | 1.77 | 0.86 | 3.24 | 0.4 | 0.49 | 0.19 | 3.26
| Analytical results for Ni in OREAS 550b (Certified Value 67 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 66.2 | 58.9 | 67 | 67 | 64 | 66 | 68.8 | 69.0723 | 63.6 | 66.6 | 69.5 | 67.6 | 67.3 | 62.5522 | 66.9 | 66 | 69.9 | 64.1 | 65.6 | 62 | 66.5 | 68 |  | 1
| 1 | 2 | 66.8 | 116.8 | 69 | 67 | 65.58 | 67 | 65.5 | 71.5542 | 66.1 | 66.6 | 68.4 | 68.2 | 69 | 63.8933 | 67.5 | 68 | 69.2 | 68 | 68.1 | 61.9 | 65.7 | 69 |  | 28
| 1 | 3 | 67.6 | 60.4 | 66 | 66 | 65.57 | 66 | 68.1 | 68.1451 | 66.4 | 65.7 | 68.3 | 73.5 | 68.4 | 64.035 | 69.1 | 68 | 69.4 | 68.1 | 66.8 | 62.3 | 64.2 | 66 |  | 16
| 1 | 4 | 67.2 | 59.9 | 66 | 66 | 64.77 | 68 | 68.6 | 72.514 | 64.1 | 68.1 | 68.6 | 71.1 | 67.6 | 64.9297 | 68.8 | 68 | 70.2 | 67.2 | 64.7 | 63.1 | 64.3 | 66 |  | 67.2048
| 1 | 5 | 64.5 | 60.4 | 66 | 66 | 66.41 | 65 | 65.8 | 70.3085 | 65.7 | 67.9 | 70.3 | 71.5 | 68.6 | 64.7307 | 67.6 | 68 | 69.4 | 66.7 | 67.1 | 62.7 | 63.9 | 68 |  | 39
| 1 | 6 | 66.8 | 58.2 | 68 | 66 | 64.46 | 66 | 67.4 | 71.06 | 66.1 | 65.9 | 67.4 | 71.1 | 67.3 | 66.2916 | 68.3 | 68 | 70.8 | 66.3 | 69.2 | 62.9 | 64.9 | 69
| Mean |  | 66.5167 | 69.1 | 67 | 66.3333 | 65.1317 | 66.3333 | 67.3667 | 70.4423 | 65.3333 | 66.8 | 68.75 | 70.5 | 68.0333 | 64.4054 | 68.0333 | 67.6667 | 69.8167 | 66.7333 | 66.9167 | 62.4833 | 64.9167 | 67.6667
| Median |  | 66.8 | 60.15 | 66.5 | 66 | 65.17 | 66 | 67.75 | 70.6842 | 65.9 | 66.6 | 68.5 | 71.1 | 68 | 64.3828 | 67.95 | 68 | 69.65 | 66.95 | 66.95 | 62.5 | 64.6 | 68
| Std Dev. |  | 1.0926 | 23.3844 | 1.2649 | 0.5164 | 0.88203 | 1.0328 | 1.418 | 1.6175 | 1.181 | 1 | 1.0134 | 2.21 | 0.72847 | 1.2468 | 0.84301 | 0.8165 | 0.608 | 1.4706 | 1.6315 | 0.4916 | 1.0048 | 1.3663
| Rel.Std.Dev. |  | 1.643% | 33.84% | 1.888% | 0.7785% | 1.354% | 1.557% | 2.105% | 2.296% | 1.808% | 1.497% | 1.474% | 3.135% | 1.071% | 1.936% | 1.239% | 1.207% | 0.8709% | 2.204% | 2.438% | 0.7868% | 1.548% | 2.019%
| PDM3 |  | -1.024% | 2.82% | -0.3047% | -1.297% | -3.085% | -1.297% | 0.2408% | 4.817% | -2.785% | -0.6023% | 2.299% | 4.903% | 1.233% | -4.165% | 1.233% | 0.6872% | 3.886% | -0.7015% | -0.4287% | -7.025% | -3.405% | 0.6872%
| Z-Score (Absolute) |  | 0.28 | 1.34 | 0.03 | 0.39 | 1.15 | 0.39 | 0.26 | 2.19 | 1.02 | 0.1 | 1.12 | 2.22 | 0.67 | 1.6 | 0.67 | 0.44 | 1.79 | 0.14 | 0.03 | 2.81 | 1.28 | 0.44
| Analytical results for P in OREAS 550b (Certified Value 0.05 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.049 | 0.048 | 0.055 | 0.052 | 0.064 | 0.0458 | 0.0465 | 0.051915 | 0.049 | 0.049 | 0.046 | 0.052 | 0.05 | 0.04798 | 0.0493 | 0.049 | 0.0502 | 0.045 | 0.049 | 0.051 | 0.0512 | 0.055 |  | 1
| 1 | 2 | 0.0496 | 0.049 | 0.055 | 0.052 | 0.056 | 0.0483 | 0.0478 | 0.051154 | 0.05 | 0.048 | 0.048 | 0.049 | 0.051 | 0.050009 | 0.0486 | 0.05 | 0.0485 | 0.048 | 0.052 | 0.051 | 0.0542 | 0.053 |  | #N/A
| 1 | 3 | 0.0501 | 0.049 | 0.055 | 0.051 | 0.077 | 0.0494 | 0.0463 | 0.050823 | 0.048 | 0.048 | 0.045 | 0.052 | 0.051 | 0.049576 | 0.0498 | 0.05 | 0.0491 | 0.0482 | 0.05 | 0.051 | 0.0492 | 0.054 |  | 16
| 1 | 4 | 0.0497 | 0.049 | 0.055 | 0.051 | 0.063 | 0.0487 | 0.0469 | 0.051778 | 0.049 | 0.049 | 0.045 | 0.051 | 0.051 | 0.050063 | 0.0491 | 0.05 | 0.0508 | 0.0479 | 0.05 | 0.052 | 0.0495 | 0.052 |  | 0.049977
| 1 | 5 | 0.0485 | 0.048 | 0.054 | 0.051 | 0.067 | 0.0497 | 0.0478 | 0.051607 | 0.047 | 0.049 | 0.046 | 0.052 | 0.051 | 0.050773 | 0.0496 | 0.05 | 0.0486 | 0.0468 | 0.052 | 0.051 | 0.0536 | 0.053 |  | 40
| 1 | 6 | 0.0508 | 0.047 | 0.054 | 0.051 | 0.068 | 0.0473 | 0.0471 | 0.050725 | 0.049 | 0.048 | 0.046 | 0.053 | 0.051 | 0.052156 | 0.0492 | 0.05 | 0.0535 | 0.0468 | 0.054 | 0.051 | 0.0505 | 0.053
| Mean |  | 0.049617 | 0.048333 | 0.054667 | 0.051333 | 0.065833 | 0.0482 | 0.047067 | 0.051334 | 0.048667 | 0.0485 | 0.046 | 0.0515 | 0.050833 | 0.050093 | 0.049267 | 0.049833 | 0.050117 | 0.047117 | 0.051167 | 0.051167 | 0.051367 | 0.053333
| Median |  | 0.04965 | 0.0485 | 0.055 | 0.051 | 0.0655 | 0.0485 | 0.047 | 0.051381 | 0.049 | 0.0485 | 0.046 | 0.052 | 0.051 | 0.050036 | 0.04925 | 0.05 | 0.04965 | 0.04735 | 0.051 | 0.051 | 0.05085 | 0.053
| Std Dev. |  | 0.0008085 | 0.0008165 | 0.0005164 | 0.0005164 | 0.0069113 | 0.0014505 | 0.00063456 | 0.00050478 | 0.0010328 | 0.00054772 | 0.0010954 | 0.0013784 | 0.00040825 | 0.0013755 | 0.00041793 | 0.00040825 | 0.0018904 | 0.001204 | 0.0018348 | 0.00040825 | 0.0020963 | 0.0010328
| Rel.Std.Dev. |  | 1.629% | 1.689% | 0.9446% | 1.006% | 10.5% | 3.009% | 1.348% | 0.9833% | 2.122% | 1.129% | 2.381% | 2.677% | 0.8031% | 2.746% | 0.8483% | 0.8192% | 3.772% | 2.555% | 3.586% | 0.7979% | 4.081% | 1.936%
| PDM3 |  | -0.7203% | -3.288% | 9.384% | 2.715% | 31.73% | -3.555% | -5.823% | 2.716% | -2.621% | -2.955% | -7.957% | 3.048% | 1.714% | 0.2323% | -1.421% | -0.2868% | 0.2801% | -5.723% | 2.381% | 2.381% | 2.781% | 6.716%
| Z-Score (Absolute) |  | 0.25 | 0.9 | 2.32 | 0.62 | 7.98 | 0.97 | 1.54 | 0.62 | 0.73 | 0.81 | 2.08 | 0.71 | 0.37 | 0.01 | 0.43 | 0.14 | 0.01 | 1.52 | 0.54 | 0.54 | 0.64 | 1.64
| Analytical results for Pb in OREAS 550b (Certified Value 13.2 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
1
| 1 | 1 | 14.1 | 11.93 | 10 | 13 | 14 | 13 | 14.1611 | 12.9 | 14.2 | 13.6 | 14 | 12.6 | 14.7617 | 13.2 | 13.9 | 15.2 | 23 | 13.1 | 10 | 12.2 | 10.7 |  |  | 1
| 1 | 2 | 15.6 | 12.67 | 9 | 13 | 15 | 12.7 | 14.3073 | 13.1 | 13.8 | 12.7 | 13.3 | 12.2 | 14.7864 | 12.9 | 14.6 | 15.9 | 15 | 12.9 | 9.8 | 12.4 | 10.6 |  |  | 13
| 1 | 3 | 14.6 | 11.27 | 9 | 13 | 14 | 12.5 | 13.8613 | 12.9 | 13.5 | 12.8 | 14.3 | 12.2 | 14.2972 | 12.3 | 14.1 | 15.3 | 13 | 13.6 | 9.7 | 12.1 | 11.6 |  |  | 16
| 1 | 4 | 16.2 | 11.45 | 8 | 13 | 14 | 13.1 | 14.512 | 12.8 | 13.6 | 13 | 14 | 12.3 | 14.4514 | 12.9 | 13.9 | 15 | 14 | 12.7 | 9.2 | 12.1 | 10.8 |  |  | 13.1704
| 1 | 5 | 14.6 | 11.39 | 9 | 13 | 14 | 12.4 | 13.8299 | 12.9 | 13.4 | 12.8 | 13.9 | 12.6 | 13.9442 | 12.6 | 14 | 15.1 | 13 | 13.1 | 11.2 | 12.5 | 11.4 |  |  | 41
| 1 | 6 | 14.5 | 10.66 | 9 | 13 | 14 | 12.8 | 14.3253 | 13.5 | 13.5 | 12.8 | 14.2 | 12.4 | 14.5428 | 12.8 | 13.7 | 15.9 | 13 | 13.2 | 11.1 | 12.7 | 10.8
| Mean |  | 14.9333 | 11.5617 | 9 | 13 | 14.1667 | 12.75 | 14.1661 | 13.0167 | 13.6667 | 12.95 | 13.95 | 12.3833 | 14.464 | 12.7833 | 14.0333 | 15.4 | 15.1667 | 13.1 | 10.1667 | 12.3333 | 10.9833
| Median |  | 14.6 | 11.42 | 9 | 13 | 14 | 12.75 | 14.2342 | 12.9 | 13.55 | 12.8 | 14 | 12.35 | 14.4971 | 12.85 | 13.95 | 15.25 | 13.5 | 13.1 | 9.9 | 12.3 | 10.8
| Std Dev. |  | 0.79415 | 0.67883 | 0.63246 | 0 | 0.40825 | 0.27386 | 0.27236 | 0.25626 | 0.29439 | 0.33317 | 0.35071 | 0.18348 | 0.31526 | 0.30605 | 0.30768 | 0.4 | 3.92 | 0.30332 | 0.80664 | 0.24221 | 0.41191
| Rel.Std.Dev. |  | 5.318% | 5.871% | 7.027% | 0% | 2.882% | 2.148% | 1.923% | 1.969% | 2.154% | 2.573% | 2.514% | 1.482% | 2.18% | 2.394% | 2.192% | 2.597% | 25.85% | 2.315% | 7.934% | 1.964% | 3.75%
| PDM3 |  | 13.39% | -12.21% | -31.67% | -1.294% | 7.564% | -3.192% | 7.56% | -1.167% | 3.768% | -1.674% | 5.919% | -5.976% | 9.821% | -2.939% | 6.552% | 16.93% | 15.16% | -0.5347% | -22.81% | -6.356% | -16.61%
| Z-Score (Absolute) |  | 1.27 | 0.97 | 2.66 | 0.01 | 0.76 | 0.18 | 0.76 | 0 | 0.43 | 0.04 | 0.62 | 0.42 | 0.96 | 0.15 | 0.67 | 1.58 | 1.43 | 0.06 | 1.89 | 0.45 | 1.35
| Analytical results for Pr in OREAS 550b (Certified Value 13.8 ppm) | [24 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 1
2
| 1 | 1 | 13.57 | 13 | 15.1 | 17.236 | 12.56 | 13.9688 | 13.9 | 9.4 | [15 empty] | 1
| 1 | 2 | 13.4 | 14.3 | 14.8 | 17.3672 | 12.86 | 14.1258 | 14.1 | 9.79 | [15 empty] | 14
| 1 | 3 | 13.35 | 12.9 | 15.1 | 17.328 | 14.17 | 13.8973 | 14.2 | 9.82 | [15 empty] | 16
| 1 | 4 | 13.5 | 12.9 | 14.6 | 17.348 | 12.7 | 14.5486 | 13.7 | 9.69 | [15 empty] | 13.7721
| 1 | 5 | 13.06 | 12.9 | 14.2 | 17.3024 | 14.26 | 14.1983 | 13.8 | 9.79 | [15 empty] | 42
| 1 | 6 | 13.72 | 11.9 | 15 | 17.3832 | 13.37 | 14.3365 | 13.8 | 9.67
| Mean |  | 13.4333 | 12.9833 | 14.8 | 17.3275 | 13.32 | 14.1792 | 13.9167 | 9.6933
| Median |  | 13.45 | 12.9 | 14.9 | 17.338 | 13.115 | 14.1621 | 13.85 | 9.74
| Std Dev. |  | 0.22483 | 0.76529 | 0.35214 | 0.053106 | 0.74595 | 0.24006 | 0.19408 | 0.15578
| Rel.Std.Dev. |  | 1.674% | 5.894% | 2.379% | 0.3065% | 5.6% | 1.693% | 1.395% | 1.607%
| PDM3 |  | -2.46% | -5.727% | 7.464% | 25.82% | -3.283% | 2.956% | 1.05% | -29.62%
| Z-Score (Absolute) |  | 0.27 | 0.78 | 1.27 | 4.12 | 0.4 | 0.57 | 0.27 | 4.49
| Analytical results for Pt in OREAS 550b (Indicative Value 7.83 ppb) | [24 empty] | Indicative
Pt | Round | Replicate | Lab | [22 empty] | 1
| No | No | 25 | [22 empty] | ppb
|  |  | 4A*OES/MS | [22 empty] | 2
2
| 1 | 1 | 9 | [22 empty] | 1
| 1 | 2 | 8 | [22 empty] | 2
| 1 | 3 | 9 | [22 empty] | 16
| 1 | 4 | 7 | [22 empty] | 7.8333
| 1 | 5 | 7 | [22 empty] | 8
| 1 | 6 | 7
| Mean |  | 7.8333
| Median |  | 7.5
| Std Dev. |  | 0.98319
| Rel.Std.Dev. |  | 12.55%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 550b (Certified Value 88 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 0
1
| 1 | 1 | 86.02 | 76.1 | 87.4 | 81.1 | 127.6275 | 89.5 | 84 | 89.2 | 87.7971 | 91.2 | 81.6 | 90.2 | 92.3 | 90.6 | 91.5533 | 84.5 | 86.6 | 90.3 | 63.6 | 82.9 | 89 | 89.5 | 53.9 | 1
| 1 | 2 | 86.93 | 94.6 | 85.3 | 80.1 | 128.5285 | 88.6 | 97 | 88.2 | 91.371 | 92.9 | 80.3 | 88.2 | 83.4 | 92.6 | 94.4361 | 83.7 | 85.9 | 91.32 | 64.2 | 83.6 | 91.8 | 86 | 61.5 | 15
| 1 | 3 | 87.11 | 74 | 88.3 | 82 | 127.413 | 92.5 | 87 | 89 | 89.8396 | 90.4 | 79.8 | 89.7 | 94.7 | 94.4 | 94.2326 | 86.2 | 85.6 | 89.93 | 65.6 | 83.9 | 89.5 | 86.3 | 60.1 | 16
| 1 | 4 | 87.05 | 82 | 86.5 | 81.8 | 127.0775 | 89 | 93 | 88.3 | 91.1558 | 90.3 | 81.5 | 89 | 89.6 | 93.7 | 97.0497 | 85.9 | 87.5 | 91.31 | 67.7 | 82.2 | 90.3 | 84.6 | 60.9 | 88.2447
| 1 | 5 | 88.46 | 85.1 | 84.9 | 82.6 | 127.262 | 94.1 | 94 | 87.5 | 91.0634 | 83.2 | 81.1 | 92.3 | 91.7 | 94 | 95.5151 | 88.2 | 89 | 89.48 | 65.6 | 85.6 | 93.9 | 84.3 | 61.5 | 43
| 1 | 6 | 87.19 | 73.4 | 88 | 80.4 | 127.6525 | 92.8 | 91 | 89.7 | 88.5024 | 93.9 | 59.6 | 88.3 | 90.7 | 92.7 | 98.7476 | 87.3 | 85.6 | 91.84 | 65.6 | 85.9 | 93.9 | 87.3 | 58.5
| Mean |  | 87.1267 | 80.8667 | 86.7333 | 81.3333 | 127.5935 | 91.0833 | 91 | 88.65 | 89.9549 | 90.3167 | 77.3167 | 89.6167 | 90.4 | 93 | 95.2557 | 85.9667 | 86.7 | 90.6967 | 65.3833 | 84.0167 | 91.4 | 86.3333 | 59.4
| Median |  | 87.08 | 79.05 | 86.95 | 81.45 | 127.5203 | 91 | 92 | 88.65 | 90.4515 | 90.8 | 80.7 | 89.35 | 91.2 | 93.2 | 94.9756 | 86.05 | 86.25 | 90.805 | 65.6 | 83.75 | 91.05 | 86.15 | 60.5
| Std Dev. |  | 0.7812 | 8.1662 | 1.4123 | 0.97091 | 0.50746 | 2.3267 | 4.7749 | 0.79687 | 1.5136 | 3.7669 | 8.7076 | 1.5276 | 3.8346 | 1.3755 | 2.4862 | 1.6801 | 1.3416 | 0.92707 | 1.4204 | 1.4689 | 2.1559 | 1.9086 | 2.9196
| Rel.Std.Dev. |  | 0.8966% | 10.1% | 1.628% | 1.194% | 0.3977% | 2.555% | 5.247% | 0.8989% | 1.683% | 4.171% | 11.26% | 1.705% | 4.242% | 1.479% | 2.61% | 1.954% | 1.547% | 1.022% | 2.172% | 1.748% | 2.359% | 2.211% | 4.915%
| PDM3 |  | -1.267% | -8.361% | -1.713% | -7.832% | 44.59% | 3.217% | 3.122% | 0.4593% | 1.938% | 2.348% | -12.38% | 1.555% | 2.442% | 5.389% | 7.945% | -2.581% | -1.75% | 2.779% | -25.91% | -4.791% | 3.576% | -2.166% | -32.69%
| Z-Score (Absolute) |  | 0.42 | 2.16 | 0.53 | 2.03 | 10.79 | 0.67 | 0.65 | 0 | 0.36 | 0.46 | 3.14 | 0.27 | 0.48 | 1.21 | 1.83 | 0.74 | 0.54 | 0.57 | 6.45 | 1.28 | 0.76 | 0.64 | 8.11
| Analytical results for Re in OREAS 550b (Certified Value 0.014 ppm) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 25 | 26 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.012 | 0.01 | 0.013 | < 0.05 | 0.041 | < 0.002 | < 0.05 | 0.012 | 0.013 | 0.011 | 0.014 | 0.013 | < 0.01 | 0.014 | < 0.003 | 0.013 | 0.018 | 0.0136 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 0.013 | 0.015 | 0.013 | < 0.05 | 0.041 | 0.014 | < 0.05 | 0.012 | 0.014 | 0.013 | 0.014 | 0.014 | < 0.01 | 0.012 | 0.007 | 0.011 | 0.019 | 0.0134 | < 0.1 |  |  |  |  | 16
| 1 | 3 | 0.013 | 0.019 | 0.013 | < 0.05 | 0.042 | 0.014 | < 0.05 | 0.015 | 0.008 | 0.014 | 0.015 | 0.013 | < 0.01 | 0.012 | < 0.003 | 0.013 | 0.018 | 0.0151 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 0.014 | 0.014 | 0.013 | < 0.05 | 0.046 | 0.035 | < 0.05 | 0.015 | 0.014 | 0.015 | 0.014 | 0.015 | < 0.01 | 0.014 | < 0.003 | 0.013 | 0.02 | 0.013 | < 0.1 |  |  |  |  | 0.013526
| 1 | 5 | 0.013 | 0.018 | 0.014 | < 0.05 | 0.04 | 0.007 | < 0.05 | 0.014 | 0.017 | 0.014 | 0.014 | 0.015 | < 0.01 | 0.015 | < 0.003 | 0.013 | 0.016 | 0.0124 | < 0.1 |  |  |  |  | 44
| 1 | 6 | 0.013 | 0.014 | 0.014 | < 0.05 | 0.041 | 0.017 | < 0.05 | 0.013 | 0.011 | 0.011 | 0.013 | 0.012 | < 0.01 | 0.014 | 0.003 | 0.013 | 0.019 | 0.0142 | < 0.1
| Mean |  | 0.013 | 0.015 | 0.013333 |  | 0.041833 | 0.0174 |  | 0.0135 | 0.012833 | 0.013 | 0.014 | 0.013667 |  | 0.0135 | 0.005 | 0.012667 | 0.018333 | 0.013617
| Median |  | 0.013 | 0.0145 | 0.013 |  | 0.041 | 0.014 |  | 0.0135 | 0.0135 | 0.0135 | 0.014 | 0.0135 |  | 0.014 | 0.005 | 0.013 | 0.0185 | 0.0135
| Std Dev. |  | 0.00063246 | 0.0032249 | 0.0005164 |  | 0.002137 | 0.010502 |  | 0.0013784 | 0.0030605 | 0.0016733 | 0.00063246 | 0.0012111 |  | 0.0012247 | 0.0028284 | 0.0008165 | 0.0013663 | 0.00094322
| Rel.Std.Dev. |  | 4.865% | 21.5% | 3.873% |  | 5.108% | 60.36% |  | 10.21% | 23.85% | 12.87% | 4.518% | 8.861% |  | 9.072% | 56.57% | 6.446% | 7.452% | 6.927%
| PDM3 |  | -3.892% | 10.89% | -1.427% |  | 209.3% | 28.64% |  | -0.1951% | -5.124% | -3.892% | 3.501% | 1.037% |  | -0.1951% | -63.04% | -6.356% | 35.54% | 0.6674%
| Z-Score (Absolute) |  | 0.44 | 0.98 | 0.2 | 8.08 | 20.03 | 0.75 | 8.08 | 0.08 | 0.56 | 0.44 | 0.27 | 0.04 | 6.12 | 0.08 | 7.77 | 0.67 | 3.35 | 0 | 25.82
| Analytical results for S in OREAS 550b (Certified Value 1.16 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.1852 | 1.12 | 1.17 | 1.24 | 1.0212 | 1.001 | 1.15 | 1.21 | 1.1894 | 1.15 | 1.18 | 1.16 | 1.15 | 1.17 | 1.1669 | 1.1 | 1.16 | 1.13 | 1.0314 | 1.1 | 1.21 | 1.19 | 1.285 | 1
| 1 | 2 | 1.1718 | 1.12 | 1.16 | 1.21 | 1.0255 | 1.074 | 1.15 | 1.16 | 1.1855 | 1.17 | 1.16 | 1.15 | 1.14 | 1.18 | 1.1667 | 1.1 | 1.17 | 1.13 | 1.1084 | 1.1 | 1.22 | 1.2 | 1.259 | 30
| 1 | 3 | 1.1808 | 1.12 | 1.16 | 1.21 | 1.0209 | 1.137 | 1.15 | 1.22 | 1.1892 | 1.15 | 1.17 | 1.1 | 1.2 | 1.18 | 1.1688 | 1.11 | 1.16 | 1.13 | 1.1156 | 1.1 | 1.22 | 1.19 | 1.25 | 16
| 1 | 4 | 1.1821 | 1.12 | 1.16 | 1.23 | 1.0219 | 1.112 | 1.16 | 1.15 | 1.1851 | 1.16 | 1.17 | 1.07 | 1.2 | 1.18 | 1.1799 | 1.13 | 1.16 | 1.16 | 1.089 | 1.1 | 1.24 | 1.17 | 1.279 | 1.1574
| 1 | 5 | 1.1385 | 1.14 | 1.14 | 1.23 | 1.028 | 1.188 | 1.16 | 1.19 | 1.185 | 1.13 | 1.15 | 1.1 | 1.2 | 1.19 | 1.136 | 1.13 | 1.18 | 1.12 | 1.077 | 1.1 | 1.22 | 1.18 | 1.268 | 45
| 1 | 6 | 1.2104 | 1.08 | 1.15 | 1.2 | 1.0177 | 1.217 | 1.15 | 1.2 | 1.1843 | 1.17 | 1.16 | 1.13 | 1.22 | 1.18 | 1.1655 | 1.11 | 1.14 | 1.11 | 1.0824 | 1.1 | 1.22 | 1.16 | 1.273
| Mean |  | 1.1781 | 1.1167 | 1.1567 | 1.22 | 1.0225 | 1.1215 | 1.1533 | 1.1883 | 1.1864 | 1.155 | 1.165 | 1.1183 | 1.185 | 1.18 | 1.164 | 1.1133 | 1.1617 | 1.13 | 1.084 | 1.1 | 1.2217 | 1.1817 | 1.269
| Median |  | 1.1814 | 1.12 | 1.16 | 1.22 | 1.0215 | 1.1245 | 1.15 | 1.195 | 1.1853 | 1.155 | 1.165 | 1.115 | 1.2 | 1.18 | 1.1668 | 1.11 | 1.16 | 1.13 | 1.0857 | 1.1 | 1.22 | 1.185 | 1.2705
| Std Dev. |  | 0.023345 | 0.019664 | 0.010328 | 0.015492 | 0.0036494 | 0.078294 | 0.005164 | 0.027869 | 0.0022373 | 0.015166 | 0.010488 | 0.034303 | 0.032094 | 0.0063246 | 0.014682 | 0.013663 | 0.013292 | 0.016733 | 0.02979 | 2.4324e-16 | 0.0098319 | 0.01472 | 0.012915
| Rel.Std.Dev. |  | 1.981% | 1.761% | 0.8929% | 1.27% | 0.3569% | 6.981% | 0.4477% | 2.345% | 0.1886% | 1.313% | 0.9003% | 3.067% | 2.708% | 0.536% | 1.261% | 1.227% | 1.144% | 1.481% | 2.748% | 2.211e-14% | 0.8048% | 1.246% | 1.018%
| PDM3 |  | 1.794% | -3.517% | -0.06069% | 5.411% | -11.65% | -3.099% | -0.3487% | 2.675% | 2.51% | -0.2047% | 0.6593% | -3.373% | 2.387% | 1.955% | 0.5696% | -3.805% | 0.3713% | -2.365% | -6.342% | -4.957% | 5.555% | 2.099% | 9.645%
| Z-Score (Absolute) |  | 0.42 | 1.14 | 0.13 | 1.48 | 3.52 | 1.02 | 0.21 | 0.67 | 0.63 | 0.17 | 0.08 | 1.1 | 0.59 | 0.46 | 0.06 | 1.22 | 0 | 0.8 | 1.96 | 1.56 | 1.52 | 0.51 | 2.71
| Analytical results for Sb in OREAS 550b (Certified Value 2.62 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.88 | 2.79 | 2.7 | 2.7 | 2.19 | 2.41 | 2.8 | 1.535 | 2.74 | 2.87 | 2.79 | 2.29 | 2.63 | 2.7686 | 2.56 | 2.76 | 2.48 | 2.8 | 2.4 | 2.6 | 2.7 | 2.2 |  | 1
| 1 | 2 | 2.66 | 3.12 | 3 | 2.7 | 2.22 | 2.56 | 2.72 | 1.6522 | 2.91 | 2.67 | 2.68 | 2.41 | 2.65 | 2.7225 | 2.45 | 2.55 | 2.54 | 2.79 | 2.3 | 2.6 | 2.8 | 2.4 |  | 17
| 1 | 3 | 2.87 | 2.82 | 3 | 2.5 | 2.49 | 2.79 | 2.89 | 1.4768 | 2.93 | 2.73 | 2.69 | 2.5 | 2.67 | 2.6838 | 2.45 | 2.56 | 2.35 | 2.83 | 2.4 | 3 | 2.7 | 2.1 |  | 16
| 1 | 4 | 2.83 | 2.92 | 2.7 | 2.5 | 2.06 | 2.65 | 2.65 | 1.5827 | 2.76 | 2.77 | 2.74 | 2.49 | 2.57 | 2.9147 | 2.39 | 2.65 | 2.38 | 2.49 | 2.2 | 2.3 | 2.7 | 2.27 |  | 2.6159
| 1 | 5 | 2.73 | 2.74 | 3 | 2.6 | 2.33 | 2.34 | 2.88 | 1.246 | 2.85 | 2.8 | 2.72 | 2.4 | 2.64 | 2.8257 | 2.49 | 2.53 | 2.27 | 2.79 | 2.2 | 2.8 | 2.7 | 2.09 |  | 46
| 1 | 6 | 2.79 | 2.55 | 2.8 | 2.4 | 2.19 | 2.59 | 2.73 | 1.3243 | 2.93 | 2.85 | 2.63 | 2.49 | 2.59 | 2.8338 | 2.36 | 2.51 | 2.39 | 2.89 | 2.2 | 2.8 | 2.7 | 2.26
| Mean |  | 2.7933 | 2.8233 | 2.8667 | 2.5667 | 2.2467 | 2.5567 | 2.7783 | 1.4695 | 2.8533 | 2.7817 | 2.7083 | 2.43 | 2.625 | 2.7915 | 2.45 | 2.5933 | 2.4017 | 2.765 | 2.2833 | 2.6833 | 2.7167 | 2.22
| Median |  | 2.81 | 2.805 | 2.9 | 2.55 | 2.205 | 2.575 | 2.765 | 1.5059 | 2.88 | 2.785 | 2.705 | 2.45 | 2.635 | 2.7972 | 2.45 | 2.555 | 2.385 | 2.795 | 2.25 | 2.7 | 2.7 | 2.23
| Std Dev. |  | 0.085479 | 0.18981 | 0.15055 | 0.12111 | 0.14706 | 0.16293 | 0.095376 | 0.15592 | 0.085479 | 0.074944 | 0.054924 | 0.08124 | 0.037815 | 0.083701 | 0.071274 | 0.094798 | 0.095795 | 0.13996 | 0.098319 | 0.24014 | 0.040825 | 0.11679
| Rel.Std.Dev. |  | 3.06% | 6.723% | 5.252% | 4.718% | 6.546% | 6.373% | 3.433% | 10.61% | 2.996% | 2.694% | 2.028% | 3.343% | 1.441% | 2.998% | 2.909% | 3.655% | 3.989% | 5.062% | 4.306% | 8.949% | 1.503% | 5.261%
| PDM3 |  | 6.781% | 7.928% | 9.584% | -1.884% | -14.12% | -2.266% | 6.208% | -43.83% | 9.075% | 6.335% | 3.532% | -7.108% | 0.3461% | 6.712% | -6.344% | -0.8644% | -8.191% | 5.698% | -12.71% | 2.576% | 3.85% | -15.14%
| Z-Score (Absolute) |  | 0.68 | 0.83 | 1.04 | 0.43 | 1.99 | 0.48 | 0.61 | 5.78 | 0.97 | 0.62 | 0.26 | 1.09 | 0.14 | 0.67 | 1 | 0.3 | 1.23 | 0.54 | 1.81 | 0.14 | 0.3 | 2.12
| Analytical results for Sc in OREAS 550b (Certified Value 11.3 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 11.1 | 11.1 | 12 | 12 | 19.5538 | 8.7 | 13.2 | 10.8 | 11.3008 | 10.2 | 11 | 10.5 | 12.1 | 11.1 | 11.1855 | 10.9 | 11.6 | 11.3 | 10 | 10 | 11.6 | 12 | 11.7 | 1
| 1 | 2 | 11 | 11.2 | 12 | 12 | 19.5756 | 8.9 | 12.5 | 11.5 | 11.5429 | 10.3 | 11 | 10.3 | 11.2 | 11.4 | 11.2394 | 11.5 | 11.7 | 11.3 | 10.65 | 11 | 11.8 | 11 | 11.9 | 18
| 1 | 3 | 11.2 | 11.1 | 12 | 12 | 19.5602 | 9.3 | 12.8 | 10.6 | 11.2617 | 10.2 | 11 | 10.7 | 12.1 | 11.5 | 11.5273 | 12 | 11.5 | 11.3 | 10.82 | 11 | 11.8 | 11 | 11.8 | 16
| 1 | 4 | 11.1 | 11 | 12 | 12 | 19.5479 | 8.7 | 12.9 | 11.3 | 11.2765 | 10.4 | 11.2 | 10.6 | 11.8 | 11.6 | 11.5589 | 11.7 | 11.5 | 11.3 | 10.6 | 10 | 11.8 | 10 | 11.6 | 11.3035
| 1 | 5 | 10.5 | 11.2 | 11 | 12 | 19.6238 | 9.5 | 12.8 | 10.9 | 11.4771 | 10.4 | 10.8 | 11 | 12 | 11.3 | 11.4156 | 11.7 | 11.7 | 11.3 | 10.4 | 11 | 12.1 | 11 | 12 | 47
| 1 | 6 | 11.3 | 10.9 | 12 | 12 | 19.5535 | 9.2 | 11.9 | 11.5 | 11.1214 | 10.7 | 10.6 | 10.4 | 11.9 | 11.2 | 11.7367 | 11.6 | 10.9 | 11.5 | 10.46 | 11 | 11.8 | 10 | 11.9
| Mean |  | 11.0333 | 11.0833 | 11.8333 | 12 | 19.5691 | 9.05 | 12.6833 | 11.1 | 11.33 | 10.3667 | 10.9333 | 10.5833 | 11.85 | 11.35 | 11.4439 | 11.5667 | 11.4833 | 11.3333 | 10.4883 | 10.6667 | 11.8167 | 10.8333 | 11.8167
| Median |  | 11.1 | 11.1 | 12 | 12 | 19.557 | 9.05 | 12.8 | 11.1 | 11.2886 | 10.35 | 11 | 10.55 | 11.95 | 11.35 | 11.4715 | 11.65 | 11.55 | 11.3 | 10.53 | 11 | 11.8 | 11 | 11.85
| Std Dev. |  | 0.28048 | 0.1169 | 0.40825 | 0 | 0.028423 | 0.33317 | 0.4446 | 0.38471 | 0.15419 | 0.18619 | 0.20656 | 0.24833 | 0.33912 | 0.18708 | 0.20751 | 0.36697 | 0.29944 | 0.08165 | 0.28131 | 0.5164 | 0.16021 | 0.75277 | 0.1472
| Rel.Std.Dev. |  | 2.542% | 1.055% | 3.45% | 0% | 0.1452% | 3.681% | 3.505% | 3.466% | 1.361% | 1.796% | 1.889% | 2.346% | 2.862% | 1.648% | 1.813% | 3.173% | 2.608% | 0.7204% | 2.682% | 4.841% | 1.356% | 6.949% | 1.246%
| PDM3 |  | -2.39% | -1.947% | 4.688% | 6.162% | 73.12% | -19.94% | 12.21% | -1.8% | 0.235% | -8.288% | -3.275% | -6.371% | 4.835% | 0.4117% | 1.243% | 2.328% | 1.591% | 0.2642% | -7.211% | -5.634% | 4.54% | -4.159% | 4.54%
| Z-Score (Absolute) |  | 0.51 | 0.42 | NA | NA | 13.88 | 3.85 | 2.28 | 0.39 | 0.01 | 1.63 | 0.67 | 1.26 | 0.87 | 0.03 | 0.19 | 0.39 | 0.25 | 0 | 1.42 | NA | 0.81 | NA | 0.81
Results from laboratories 3, 4, 20 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 550b (Certified Value 3.69 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 3.6 | 3.9 | 3.3 | 4 | 6 | 3 | 3 | 3.1135 | 4 | 4 | 3 | 3 | 4 | 4.0619 | 3.25 | 4 | 3.3 | 2.41 | 4 | 3.77 | < 10 |  |  | 1
| 1 | 2 | 3.6 | 4 | 3.3 | 4 | 4 | 3 | 4 | 3.2512 | 3 | 3 | 4 | 4 | 4 | 4.1877 | 3.07 | 4 | 3.5 | 2.3 | 4 | 3.64 | < 10 |  |  | 22
| 1 | 3 | 3.5 | 4 | 3.4 | 4 | 3 | 3 | 2 | 3.5236 | 4 | 4 | 4 | 4 | 4 | 3.983 | 3.52 | 6 | 3.9 | 2.81 | 4 | 3.53 | < 10 |  |  | 16
| 1 | 4 | 3.5 | 4.5 | 3.3 | 4 | 4 | 3 | 3 | 3.4733 | 4 | 4 | 3 | 4 | 4 | 4.4963 | 4.03 | 4 | 3.8 | 2.13 | 4 | 3.63 | < 10 |  |  | 3.6942
| 1 | 5 | 3.2 | 4.4 | 3.2 | 4 | 5 | 2 | 4 | 3.163 | 4 | 5 | 4 | 3 | 4 | 4.0394 | 4.18 | 4 | 3.6 | 2.98 | 4 | 3.68 | < 10 |  |  | 48
| 1 | 6 | 3.8 | 3.2 | 3.4 | 4 | 4 | 3 | 2 | 3.6382 | 4 | 3 | 4 | 4 | 4 | 3.876 | 4.26 | 4 | 4.1 | 3.09 | 4 | 3.47 | < 10
| Mean |  | 3.5333 | 4 | 3.3167 | 4 | 4.3333 | 2.8333 | 3 | 3.3605 | 3.8333 | 3.8333 | 3.6667 | 3.6667 | 4 | 4.1074 | 3.7183 | 4.3333 | 3.7 | 2.62 | 4 | 3.62
| Median |  | 3.55 | 4 | 3.3 | 4 | 4 | 3 | 3 | 3.3622 | 4 | 4 | 4 | 4 | 4 | 4.0507 | 3.775 | 4 | 3.7 | 2.61 | 4 | 3.635
| Std Dev. |  | 0.19664 | 0.46043 | 0.075277 | 0 | 1.0328 | 0.40825 | 0.89443 | 0.21372 | 0.40825 | 0.75277 | 0.5164 | 0.5164 | 0 | 0.21605 | 0.50649 | 0.8165 | 0.28983 | 0.39324 | 0 | 0.10696
| Rel.Std.Dev. |  | 5.565% | 11.51% | 2.27% | 0% | 23.83% | 14.41% | 29.81% | 6.36% | 10.65% | 19.64% | 14.08% | 14.08% | 0% | 5.26% | 13.62% | 18.84% | 7.833% | 15.01% | 0% | 2.955%
| PDM3 |  | -4.354% | 8.278% | -10.22% | 8.278% | 17.3% | -23.3% | -18.79% | -9.034% | 3.767% | 3.767% | -0.745% | -0.745% | 8.278% | 11.19% | 0.6536% | 17.3% | 0.1574% | -29.08% | 8.278% | -2.008%
| Z-Score (Absolute) |  | 0.44 | 0.67 | 0.96 | 0.67 | 1.47 | 2.12 | 1.72 | 0.86 | 0.28 | 0.28 | 0.12 | 0.12 | 0.67 | 0.93 | 0 | 1.47 | 0.04 | 2.63 | 0.67 | 0.24 | 3.07
| Analytical results for Sm in OREAS 550b (Certified Value 7.4 ppm) | [24 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 7.22 | 6.5 | 8.1 | 8.5104 | 6.42 | 7.1874 | 7.49 | 5.1 | [15 empty] | 1
| 1 | 2 | 7.04 | 6.9 | 8.4 | 8.5576 | 7 | 7.2633 | 7.28 | 5.3 | [15 empty] | 5
| 1 | 3 | 7.17 | 6.4 | 7.9 | 8.5776 | 7.51 | 7.0029 | 7.31 | 5.4 | [15 empty] | 16
| 1 | 4 | 7.04 | 6.5 | 7.9 | 8.5504 | 6.87 | 7.5048 | 6.96 | 5.3 | [15 empty] | 7.3981
| 1 | 5 | 6.91 | 6.2 | 7.9 | 8.5808 | 7.66 | 7.1857 | 7.39 | 5.3 | [15 empty] | 49
| 1 | 6 | 7.48 | 5.9 | 8.3 | 8.5504 | 7.06 | 7.2294 | 7.31 | 5.3
| Mean |  | 7.1433 | 6.4 | 8.0833 | 8.5545 | 7.0867 | 7.2289 | 7.29 | 5.2833
| Median |  | 7.105 | 6.45 | 8 | 8.554 | 7.03 | 7.2084 | 7.31 | 5.3
| Std Dev. |  | 0.19786 | 0.33466 | 0.22286 | 0.025351 | 0.44889 | 0.16246 | 0.17877 | 0.098319
| Rel.Std.Dev. |  | 2.77% | 5.229% | 2.757% | 0.2963% | 6.334% | 2.247% | 2.452% | 1.861%
| PDM3 |  | -3.444% | -13.49% | 9.262% | 15.63% | -4.21% | -2.287% | -1.461% | -28.59%
| Z-Score (Absolute) |  | 0.06 | 1.19 | 1.36 | 2.07 | 0.15 | 0.06 | 0.16 | 2.88
| Analytical results for Sn in OREAS 550b (Certified Value 2.47 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.5 | 2.7 | 3 | 2.5 | 2.8 | 2.1 | 2.5 | 2.5235 | 2.8 | 2.6 | 2.5 | 2.6 | 2.3 | 2.5311 | 2.49 | 2.4 | 2.5 | 2.21 | 2.2 | 2.5 | 2.5 | 2.2 |  | 1
| 1 | 2 | 2.6 | 2.9 | 3 | 2.5 | 2.4 | 2.2 | 2.3 | 2.5268 | 2.7 | 2.6 | 2.6 | 2.4 | 2.4 | 2.7236 | 2.57 | 2.4 | 2.5 | 2.9 | 2.2 | 2.6 | 2.4 | 2.2 |  | 6
| 1 | 3 | 2.5 | 2.6 | 3 | 2.4 | 2.6 | 2.3 | 2.5 | 2.2314 | 2.7 | 2.5 | 2.5 | 2.5 | 2.4 | 2.6279 | 2.53 | 2.4 | 2.5 | 3.4 | 2.3 | 2.7 | 2.6 | 2.1 |  | 16
| 1 | 4 | 2.5 | 2.7 | 3 | 2.6 | 2.2 | 2.2 | 2.4 | 2.5688 | 2.7 | 2.5 | 2.5 | 2.6 | 2.4 | 2.6369 | 2.5 | 2.4 | 2.6 | 2.41 | 2.2 | 2.7 | 2.5 | 2.2 |  | 2.4713
| 1 | 5 | 2.5 | 2.6 | 3 | 2.5 | 2.3 | 2.3 | 2.3 | 2.186 | 2.6 | 2.6 | 2.5 | 2.5 | 2.4 | 2.6506 | 2.59 | 2.5 | 2.4 | 2.81 | 2.2 | 2.5 | 2.4 | 2.1 |  | 50
| 1 | 6 | 2.6 | 2.4 | 3 | 2.5 | 2.2 | 2.2 | 2.5 | 2.5259 | 2.8 | 2.5 | 2.5 | 2.5 | 2.4 | 2.6966 | 2.47 | 2.4 | 2.6 | 2.51 | 2.2 | 2.3 | 2.4 | 2.1
| Mean |  | 2.5333 | 2.65 | 3 | 2.5 | 2.4167 | 2.2167 | 2.4167 | 2.4271 | 2.7167 | 2.55 | 2.5167 | 2.5167 | 2.3833 | 2.6444 | 2.525 | 2.4167 | 2.5167 | 2.7067 | 2.2167 | 2.55 | 2.4667 | 2.15
| Median |  | 2.5 | 2.65 | 3 | 2.5 | 2.35 | 2.2 | 2.45 | 2.5247 | 2.7 | 2.55 | 2.5 | 2.5 | 2.4 | 2.6437 | 2.515 | 2.4 | 2.5 | 2.66 | 2.2 | 2.55 | 2.45 | 2.15
| Std Dev. |  | 0.05164 | 0.16432 | 0 | 0.063246 | 0.24014 | 0.075277 | 0.098319 | 0.17058 | 0.075277 | 0.054772 | 0.040825 | 0.075277 | 0.040825 | 0.066657 | 0.047223 | 0.040825 | 0.075277 | 0.42477 | 0.040825 | 0.15166 | 0.08165 | 0.054772
| Rel.Std.Dev. |  | 2.038% | 6.201% | 0% | 2.53% | 9.937% | 3.396% | 4.068% | 7.028% | 2.771% | 2.148% | 1.622% | 2.991% | 1.713% | 2.521% | 1.87% | 1.689% | 2.991% | 15.69% | 1.842% | 5.947% | 3.31% | 2.548%
| PDM3 |  | 2.509% | 7.23% | 21.39% | 1.16% | -2.212% | -10.3% | -2.212% | -1.791% | 9.927% | 3.183% | 1.835% | 1.835% | -3.561% | 7.005% | 2.172% | -2.212% | 1.835% | 9.523% | -10.3% | 3.183% | -0.1885% | -13%
| Z-Score (Absolute) |  | 0.11 | 0.9 | NA | 0.11 | 0.67 | 2.02 | 0.67 | 0.6 | 1.35 | 0.22 | 0 | 0 | 0.9 | 0.86 | 0.06 | 0.67 | 0 | 1.28 | 2.02 | 0.22 | 0.34 | 2.47
Results from laboratory 3 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 550b (Certified Value 77 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 77.59 | 79 | 79.4 | 77.4 | 76.735 | 78.9 | 77 | 75.2 | 80.6067 | 79.1 | 79.2 | 74.8 | 78.3 | 77.1 | 76.1834 | 73.4 | 74.7 | 78.41 | 81 | 72 | 78 | 69.6 | 69 | 1
| 1 | 2 | 77.18 | 79 | 77.1 | 77.8 | 74.6275 | 89.4 | 76 | 77.3 | 82.4445 | 80.6 | 79.3 | 73.8 | 76.7 | 77.5 | 76.0362 | 73.4 | 74.5 | 79.49 | 83 | 70 | 79 | 68.9 | 72 | #N/A
| 1 | 3 | 78.33 | 80 | 81.9 | 78.6 | 74.625 | 88.3 | 76 | 74.9 | 79.4322 | 79.3 | 80.4 | 73.5 | 81.8 | 78.7 | 77.5572 | 74.8 | 74.9 | 78.68 | 84 | 75 | 79 | 70.9 | 71 | 16
| 1 | 4 | 78.07 | 80 | 79.3 | 78.6 | 75.045 | 88.9 | 77 | 76.5 | 83.5246 | 79.3 | 81.7 | 72.4 | 82.4 | 78 | 78.0704 | 74 | 75 | 80.97 | 82 | 70 | 78 | 70.6 | 71 | 77.4751
| 1 | 5 | 75.5 | 81 | 78.6 | 78.3 | 76.683 | 92.2 | 79 | 77.2 | 81.9758 | 78.4 | 82.2 | 74.2 | 81.4 | 78.6 | 77.2088 | 76.9 | 76 | 78.81 | 82 | 71 | 79 | 70.9 | 72 | 51
| 1 | 6 | 79.14 | 78 | 81.1 | 78.4 | 76.3415 | 85.7 | 79 | 74.3 | 80.8294 | 81.1 | 78.5 | 72.6 | 81.4 | 79.1 | 78.9706 | 74.6 | 75 | 80.29 | 82 | 71 | 79 | 67.3 | 71
| Mean |  | 77.635 | 79.5 | 79.5667 | 78.1833 | 75.6762 | 87.2333 | 77.3333 | 75.9 | 81.4689 | 79.6333 | 80.2167 | 73.55 | 80.3333 | 78.1667 | 77.3378 | 74.5167 | 75.0167 | 79.4417 | 82.3333 | 71.5 | 78.6667 | 69.7 | 71
| Median |  | 77.83 | 79.5 | 79.35 | 78.35 | 75.6933 | 88.6 | 77 | 75.85 | 81.4026 | 79.3 | 79.85 | 73.65 | 81.4 | 78.3 | 77.383 | 74.3 | 74.95 | 79.15 | 82 | 71 | 79 | 70.1 | 71
| Std Dev. |  | 1.2413 | 1.0488 | 1.727 | 0.48339 | 1.0179 | 4.5842 | 1.3663 | 1.2696 | 1.4653 | 1.0113 | 1.4825 | 0.92466 | 2.2818 | 0.76855 | 1.122 | 1.306 | 0.51929 | 1.0101 | 1.0328 | 1.8708 | 0.5164 | 1.4213 | 1.0954
| Rel.Std.Dev. |  | 1.599% | 1.319% | 2.171% | 0.6183% | 1.345% | 5.255% | 1.767% | 1.673% | 1.799% | 1.27% | 1.848% | 1.257% | 2.84% | 0.9832% | 1.451% | 1.753% | 0.6922% | 1.271% | 1.254% | 2.617% | 0.6564% | 2.039% | 1.543%
| PDM3 |  | 0.2064% | 2.614% | 2.7% | 0.9142% | -2.322% | 12.6% | -0.1829% | -2.033% | 5.155% | 2.786% | 3.539% | -5.066% | 3.689% | 0.8927% | -0.1772% | -3.819% | -3.173% | 2.538% | 6.271% | -7.712% | 1.538% | -10.04% | -8.358%
| Z-Score (Absolute) |  | 0.17 | 0.41 | 0.44 | 0.01 | 0.78 | 2.82 | 0.26 | 0.71 | 1.03 | 0.46 | 0.64 | 1.44 | 0.67 | 0 | 0.26 | 1.14 | 0.98 | 0.4 | 1.3 | 2.07 | 0.16 | 2.63 | 2.23
| Analytical results for Ta in OREAS 550b (Certified Value 0.74 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.79 | 0.8 | 1.4 | 0.57 | 0.74 | 0.3 | 0.67 | 0.64397 | 0.98 | 1 | 0.95 | 0.68 | 0.95 | 1.6336 | 0.53 | 0.84 | 0.19 | 0.7 | 1.8 | 0.9 | 0.96 |  |  | 1
| 1 | 2 | 0.8 | 0.9 | 1.4 | 0.53 | 1 | 0.27 | 0.68 | 0.63906 | 0.96 | 0.97 | 0.97 | 0.75 | 0.93 | 1.2593 | 0.52 | 0.77 | 0.2 | 0.6 | 1.6 | 0.86 | 0.83 |  |  | 7
| 1 | 3 | 0.77 | 0.8 | 1.4 | 0.44 | 0.61 | 0.29 | 0.66 | 0.62791 | 0.98 | 0.97 | 0.95 | 0.84 | 0.95 | 1.3517 | 0.54 | 0.75 | 0.18 | 0.7 | 1.5 | 0.85 | 0.82 |  |  | 16
| 1 | 4 | 0.77 | 0.8 | 1.4 | 0.38 | 0.73 | 0.3 | 0.7 | 0.64239 | 0.94 | 0.98 | 0.98 | 0.74 | 0.94 | 1.3062 | 0.48 | 0.73 | 0.21 | 0.6 | 1.4 | 0.8 | 0.82 |  |  | 0.74095
| 1 | 5 | 0.81 | 0.8 | 1.4 | 0.35 | 0.69 | 0.3 | 0.68 | 0.62039 | 0.95 | 0.96 | 0.94 | 0.71 | 0.92 | 1.1913 | 0.5 | 0.6 | 0.17 | 0.6 | 1.4 | 0.79 | 0.77 |  |  | 52
| 1 | 6 | 0.8 | 0.8 | 1.4 | 0.42 | 0.81 | 0.32 | 0.7 | 0.62151 | 0.95 | 0.93 | 0.93 | 0.74 | 0.98 | 1.115 | 0.49 | 0.68 | 0.17 | 0.6 | 1.9 | 0.94 | 0.82
| Mean |  | 0.79 | 0.81667 | 1.4 | 0.44833 | 0.76333 | 0.29667 | 0.68167 | 0.63254 | 0.96 | 0.96833 | 0.95333 | 0.74333 | 0.945 | 1.3095 | 0.51 | 0.72833 | 0.18667 | 0.63333 | 1.6 | 0.85667 | 0.83667
| Median |  | 0.795 | 0.8 | 1.4 | 0.43 | 0.735 | 0.3 | 0.68 | 0.63349 | 0.955 | 0.97 | 0.95 | 0.74 | 0.945 | 1.2828 | 0.51 | 0.74 | 0.185 | 0.6 | 1.55 | 0.855 | 0.82
| Std Dev. |  | 0.016733 | 0.040825 | 2.4324e-16 | 0.085654 | 0.13322 | 0.01633 | 0.016021 | 0.01059 | 0.016733 | 0.023166 | 0.018619 | 0.053914 | 0.020736 | 0.17952 | 0.023664 | 0.081833 | 0.01633 | 0.05164 | 0.20976 | 0.057504 | 0.064083
| Rel.Std.Dev. |  | 2.118% | 4.999% | 1.737e-14% | 19.11% | 17.45% | 5.504% | 2.35% | 1.674% | 1.743% | 2.392% | 1.953% | 7.253% | 2.194% | 13.71% | 4.64% | 11.24% | 8.748% | 8.154% | 13.11% | 6.712% | 7.659%
| PDM3 |  | 6.62% | 10.22% | 88.95% | -39.49% | 3.021% | -59.96% | -8.001% | -14.63% | 29.56% | 30.69% | 28.66% | 0.3221% | 27.54% | 76.74% | -31.17% | -1.702% | -74.81% | -14.52% | 115.9% | 15.62% | 12.92%
| Z-Score (Absolute) |  | 0.1 | NA | NA | 1.17 | 0 | 1.73 | 0.3 | 0.49 | 0.73 | 0.76 | 0.71 | 0.07 | 0.67 | 2.03 | 0.94 | 0.13 | 2.14 | NA | NA | 0.35 | 0.27
Results from laboratories 2, 3, 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 550b (Certified Value 0.68 ppm) | [24 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 15 | 16 | 17 | 25 | 26 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.68 | 0.7 | 0.8 | 0.9556 | 0.6 | 0.68 | 0.67851 | 0.71 | 0.69 | 0.68 | 0.48 | [12 empty] | 1
| 1 | 2 | 0.63 | 0.7 | 0.8 | 0.95384 | 0.62 | 0.62 | 0.72002 | 0.74 | 0.64 | 0.71 | 0.5 | [12 empty] | 8
| 1 | 3 | 0.65 | 0.6 | 0.8 | 0.95094 | 0.67 | 0.7 | 0.66275 | 0.75 | 0.66 | 0.69 | 0.51 | [12 empty] | 16
| 1 | 4 | 0.64 | 0.6 | 0.8 | 0.9584 | 0.64 | 0.66 | 0.75781 | 0.74 | 0.66 | 0.68 | 0.5 | [12 empty] | 0.6795
| 1 | 5 | 0.62 | 0.6 | 0.8 | 0.94288 | 0.69 | 0.63 | 0.75108 | 0.74 | 0.69 | 0.67 | 0.51 | [12 empty] | 53
| 1 | 6 | 0.68 | 0.6 | 0.8 | 0.94392 | 0.65 | 0.67 | 0.69891 | 0.74 | 0.66 | 0.69 | 0.51
| Mean |  | 0.65 | 0.63333 | 0.8 | 0.95093 | 0.645 | 0.66 | 0.71152 | 0.73667 | 0.66667 | 0.68667 | 0.50167
| Median |  | 0.645 | 0.6 | 0.8 | 0.95239 | 0.645 | 0.665 | 0.70947 | 0.74 | 0.66 | 0.685 | 0.505
| Std Dev. |  | 0.025298 | 0.05164 | 1.2162e-16 | 0.0063249 | 0.032711 | 0.030332 | 0.03849 | 0.013663 | 0.019664 | 0.013663 | 0.01169
| Rel.Std.Dev. |  | 3.892% | 8.154% | 1.52e-14% | 0.6651% | 5.071% | 4.596% | 5.41% | 1.855% | 2.95% | 1.99% | 2.33%
| PDM3 |  | -4.342% | -6.795% | 17.73% | 39.95% | -5.078% | -2.87% | 4.711% | 8.413% | -1.889% | 1.054% | -26.17%
| Z-Score (Absolute) |  | 0.52 | NA | NA | 8.85 | 0.67 | 0.21 | 1.4 | 2.18 | 0 | 0.62 | 5.14
Results from laboratories 2 and 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 550b (Certified Value 0.07 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 25 | 26 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.12 | < 0.1 | < 0.1 | 0.2 | 1.91 | 0.08 | < 0.1 | 0.06 | 0.07 | 0.08 | 0.06 | 0.06 | < 0.2 | 0.05 | < 0.03 | < 0.5 | 0.1 | 0.06 | < 0.05 |  |  |  |  | 1
| 1 | 2 | 0.09 | < 0.1 | < 0.1 | 0.24 | 1.81 | 0.08 | < 0.1 | 0.07 | 0.05 | < 0.05 | 0.06 | 0.08 | < 0.2 | 0.05 | 1.69 | < 0.5 | < 0.1 | 0.05 | < 0.05 |  |  |  |  | 9
| 1 | 3 | 0.13 | < 0.1 | < 0.1 | 0.26 | 1.88 | 0.1 | < 0.1 | 0.11 | 0.06 | 0.05 | < 0.05 | 0.05 | < 0.2 | 0.06 | < 0.03 | < 0.5 | 0.1 | 0.06 | < 0.05 |  |  |  |  | 16
| 1 | 4 | < 0.05 | < 0.1 | < 0.1 | 0.19 | 1.93 | 0.07 | < 0.1 | 0.06 | 0.05 | 0.06 | 0.07 | 0.07 | < 0.2 | 0.05 | < 0.03 | < 0.5 | 0.1 | 0.06 | < 0.05 |  |  |  |  | 0.07025
| 1 | 5 | 0.06 | < 0.1 | < 0.1 | 0.19 | 1.94 | 0.09 | < 0.1 | 0.07 | 0.08 | 0.06 | 0.05 | 0.05 | < 0.2 | 0.06 | < 0.03 | < 0.5 | < 0.1 | 0.06 | < 0.05 |  |  |  |  | 54
| 1 | 6 | 0.09 | < 0.1 | < 0.1 | 0.18 | 1.84 | 0.08 | < 0.1 | 0.05 | 0.07 | < 0.05 | 0.09 | < 0.05 | < 0.2 | 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.07 | < 0.05
| Mean |  | 0.098 |  |  | 0.21 | 1.885 | 0.083333 |  | 0.07 | 0.063333 | 0.0625 | 0.066 | 0.062 |  | 0.053333 | 1.69 |  | 0.1 | 0.06
| Median |  | 0.09 |  |  | 0.195 | 1.895 | 0.08 |  | 0.065 | 0.065 | 0.06 | 0.06 | 0.06 |  | 0.05 | 1.69 |  | 0.1 | 0.06
| Std Dev. |  | 0.027749 |  |  | 0.032249 | 0.051672 | 0.010328 |  | 0.020976 | 0.012111 | 0.012583 | 0.015166 | 0.013038 |  | 0.005164 |  |  | 1.6997e-17 | 0.0063246
| Rel.Std.Dev. |  | 28.32% |  |  | 15.36% | 2.741% | 12.39% |  | 29.97% | 19.12% | 20.13% | 22.98% | 21.03% |  | 9.682% |  |  | 1.7e-14% | 10.54%
| PDM3 |  | 39.5% |  |  | 198.9% | 2583% | 18.62% |  | -0.3559% | -9.846% | -11.03% | -6.05% | -11.74% |  | -24.08% | 2306% |  | 42.35% | -14.59%
| Z-Score (Absolute) |  | 1.14 | 0.67 | 0.67 | 7.42 | 92.13 | 1.01 | 0.67 | 0.34 | 0 | 0.67 | 0.21 | 0.38 | 1.85 | 0.51 | 11.67 | 9.44 | 0.59 | 0.17 | 1.94
| Analytical results for Th in OREAS 550b (Certified Value 13.3 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 13.02 | 13.3 | 14 | 12.2 | 12.56 | 14.29 | 13.3 | 14.0419 | 13.3 | 14.05 | 13.25 | 13.3 | 12.25 | 14.1477 | 13.78 | 13.31 | 13.37 | 15.9 | 12.5 | 13.7 | 12.8 | 12.11 |  | 1
| 1 | 2 | 12.94 | 14.4 | 14.2 | 12.5 | 13.21 | 14.79 | 14.4 | 14.4041 | 13.65 | 13.9 | 13.05 | 12.25 | 12.55 | 13.6913 | 14.69 | 13.54 | 13.25 | 15.22 | 13 | 13.9 | 13.4 | 12 |  | 10
| 1 | 3 | 13.13 | 13.3 | 14.4 | 12.3 | 14.05 | 13.96 | 13.5 | 13.3139 | 13.8 | 13.8 | 12.9 | 13.15 | 12.8 | 13.3913 | 13.66 | 13.24 | 13.22 | 13.67 | 13.3 | 13.8 | 12.9 | 11.76 |  | 16
| 1 | 4 | 12.97 | 13.4 | 13.9 | 12.4 | 13.04 | 14.54 | 14 | 14.1456 | 13.4 | 13.75 | 13.1 | 12.8 | 12.4 | 13.945 | 14.05 | 12.84 | 13.32 | 17.16 | 12.6 | 14.2 | 12.8 | 11.89 |  | 13.3329
| 1 | 5 | 12.75 | 13.1 | 13.8 | 12.5 | 14.23 | 14.52 | 13.5 | 13.3957 | 13.1 | 13.7 | 12.85 | 13.25 | 12.5 | 13.7175 | 13.98 | 13.32 | 13.29 | 14.68 | 12.7 | 14.1 | 13 | 12.18 |  | 55
| 1 | 6 | 13.14 | 12.2 | 13.7 | 12.3 | 13.7 | 14.55 | 14.3 | 13.9054 | 14 | 12 | 12.95 | 13.1 | 12.7 | 13.7188 | 13.53 | 13.35 | 13.24 | 15.27 | 13.1 | 13.9 | 13.5 | 12.12
| Mean |  | 12.9917 | 13.2833 | 14 | 12.3667 | 13.465 | 14.4417 | 13.8333 | 13.8678 | 13.5417 | 13.5333 | 13.0167 | 12.975 | 12.5333 | 13.7686 | 13.9483 | 13.2667 | 13.2817 | 15.3167 | 12.8667 | 13.9333 | 13.0667 | 12.01
| Median |  | 12.995 | 13.3 | 13.95 | 12.35 | 13.455 | 14.53 | 13.75 | 13.9736 | 13.525 | 13.775 | 13 | 13.125 | 12.525 | 13.7181 | 13.88 | 13.315 | 13.27 | 15.245 | 12.85 | 13.9 | 12.95 | 12.055
| Std Dev. |  | 0.14386 | 0.70261 | 0.26077 | 0.12111 | 0.6399 | 0.28421 | 0.46332 | 0.43039 | 0.33529 | 0.76136 | 0.1472 | 0.39592 | 0.19916 | 0.25624 | 0.4118 | 0.23201 | 0.056362 | 1.1719 | 0.31411 | 0.18619 | 0.30768 | 0.16
| Rel.Std.Dev. |  | 1.107% | 5.289% | 1.863% | 0.9793% | 4.752% | 1.968% | 3.349% | 3.104% | 2.476% | 5.626% | 1.131% | 3.051% | 1.589% | 1.861% | 2.952% | 1.749% | 0.4244% | 7.651% | 2.441% | 1.336% | 2.355% | 1.332%
| PDM3 |  | -2.559% | -0.3719% | 5.003% | -7.247% | 0.9906% | 8.316% | 3.753% | 4.012% | 1.566% | 1.503% | -2.372% | -2.685% | -5.997% | 3.268% | 4.616% | -0.4969% | -0.3844% | 14.88% | -3.497% | 4.503% | -1.997% | -9.922%
| Z-Score (Absolute) |  | 0.6 | 0.14 | 0.98 | 1.58 | 0.14 | 1.68 | 0.72 | 0.78 | 0.26 | 0.25 | 0.56 | 0.63 | 1.32 | 0.62 | 0.9 | 0.17 | 0.15 | 3.05 | 0.8 | 0.88 | 0.48 | 2.14
| Analytical results for Ti in OREAS 550b (Certified Value 0.214 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.226 | 0.288 | 0.372 | 0.185 | 0.1734 | 0.113 | 0.13 | 0.16 | 0.18998 | 0.261 | 0.278 | 0.267 | 0.202 | 0.273 | 0.25385 | 0.29 | 0.189 | 0.2351 | 0.1232 | 0.22 | 0.213 | 0.244 | 0.258 | 1
| 1 | 2 | 0.223 | 0.278 | 0.371 | 0.162 | 0.17406 | 0.131 | 0.11 | 0.14 | 0.18403 | 0.266 | 0.274 | 0.274 | 0.217 | 0.275 | 0.25453 | 0.29 | 0.194 | 0.2249 | 0.1289 | 0.203 | 0.21 | 0.234 | 0.275 | 25
| 1 | 3 | 0.2279 | 0.294 | 0.368 | 0.147 | 0.17395 | 0.122 | 0.13 | 0.16 | 0.19031 | 0.265 | 0.274 | 0.259 | 0.233 | 0.273 | 0.25183 | 0.29 | 0.201 | 0.2064 | 0.1296 | 0.206 | 0.22 | 0.238 | 0.255 | 16
| 1 | 4 | 0.2228 | 0.29 | 0.37 | 0.141 | 0.1742 | 0.122 | 0.14 | 0.15 | 0.1861 | 0.265 | 0.272 | 0.254 | 0.232 | 0.277 | 0.25285 | 0.29 | 0.192 | 0.1984 | 0.1235 | 0.198 | 0.216 | 0.234 | 0.245 | 0.21423
| 1 | 5 | 0.2333 | 0.28 | 0.365 | 0.128 | 0.17392 | 0.139 | 0.14 | 0.14 | 0.18507 | 0.252 | 0.266 | 0.261 | 0.214 | 0.281 | 0.25138 | 0.29 | 0.202 | 0.1973 | 0.1273 | 0.198 | 0.214 | 0.239 | 0.245 | 56
| 1 | 6 | 0.2345 | 0.285 | 0.372 | 0.149 | 0.17452 | 0.118 | 0.14 | 0.16 | 0.1905 | 0.265 | 0.267 | 0.26 | 0.223 | 0.283 | 0.25076 | 0.29 | 0.205 | 0.1993 | 0.1228 | 0.199 | 0.206 | 0.236 | 0.245
| Mean |  | 0.22792 | 0.28583 | 0.36967 | 0.152 | 0.17401 | 0.12417 | 0.13167 | 0.15167 | 0.18767 | 0.26233 | 0.27183 | 0.2625 | 0.22017 | 0.277 | 0.25253 | 0.29 | 0.19717 | 0.21023 | 0.12588 | 0.204 | 0.21317 | 0.2375 | 0.25383
| Median |  | 0.22695 | 0.2865 | 0.3705 | 0.148 | 0.17401 | 0.122 | 0.135 | 0.155 | 0.18804 | 0.265 | 0.273 | 0.2605 | 0.22 | 0.276 | 0.25234 | 0.29 | 0.1975 | 0.20285 | 0.1254 | 0.201 | 0.2135 | 0.237 | 0.25
| Std Dev. |  | 0.0050269 | 0.00608 | 0.0027325 | 0.019596 | 0.00037145 | 0.0093684 | 0.01169 | 0.0098319 | 0.0029236 | 0.0053541 | 0.0045789 | 0.0070071 | 0.011754 | 0.0041952 | 0.0014699 | 0 | 0.0063692 | 0.015969 | 0.003076 | 0.0084617 | 0.0048339 | 0.0037815 | 0.011839
| Rel.Std.Dev. |  | 2.206% | 2.127% | 0.7392% | 12.89% | 0.2135% | 7.545% | 8.879% | 6.483% | 1.558% | 2.041% | 1.684% | 2.669% | 5.339% | 1.515% | 0.5821% | 0% | 3.23% | 7.596% | 2.444% | 4.148% | 2.268% | 1.592% | 4.664%
| PDM3 |  | 6.388% | 33.42% | 72.56% | -29.05% | -18.77% | -42.04% | -38.54% | -29.2% | -12.4% | 22.45% | 26.89% | 22.53% | 2.771% | 29.3% | 17.88% | 35.37% | -7.965% | -1.866% | -41.24% | -4.776% | -0.4967% | 10.86% | 18.49%
| Z-Score (Absolute) |  | 0.12 | 1.05 | 2.38 | 1.09 | 0.74 | 1.53 | 1.41 | 1.09 | 0.52 | 0.67 | 0.82 | 0.67 | 0 | 0.91 | 0.52 | 1.11 | 0.37 | 0.16 | 1.5 | 0.26 | 0.11 | 0.28 | 0.54
| Analytical results for Tl in OREAS 550b (Certified Value 0.54 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 0.57 | 0.48 | 0.61 | 0.53 | 0.5 | 0.41 | 0.51 | 0.593 | 0.53 | 0.55 | 0.54 | 0.56 | 0.51 | 0.59311 | 0.53 | 0.55 | 0.6 | 0.41 | 0.6 | 0.5 | 0.5 | 0.44 |  | 1
| 1 | 2 | 0.59 | 0.5 | 0.58 | 0.55 | 0.54 | 0.42 | 0.54 | 0.606 | 0.55 | 0.55 | 0.53 | 0.54 | 0.52 | 0.59733 | 0.55 | 0.58 | 0.59 | 0.422 | 0.5 | 0.6 | 0.51 | 0.44 |  | 26
| 1 | 3 | 0.58 | 0.49 | 0.59 | 0.54 | 0.6 | 0.46 | 0.51 | 0.55162 | 0.53 | 0.53 | 0.54 | 0.56 | 0.53 | 0.58042 | 0.54 | 0.55 | 0.59 | 0.398 | 0.5 | 0.6 | 0.5 | 0.43 |  | 16
| 1 | 4 | 0.58 | 0.48 | 0.58 | 0.55 | 0.54 | 0.46 | 0.54 | 0.60402 | 0.52 | 0.54 | 0.54 | 0.56 | 0.52 | 0.60591 | 0.54 | 0.54 | 0.58 | 0.706 | 0.5 | 0.5 | 0.5 | 0.45 |  | 0.53653
| 1 | 5 | 0.57 | 0.47 | 0.56 | 0.54 | 0.61 | 0.43 | 0.52 | 0.55457 | 0.52 | 0.53 | 0.54 | 0.55 | 0.53 | 0.57697 | 0.53 | 0.54 | 0.58 | 0.432 | 0.6 | 0.6 | 0.5 | 0.46 |  | 57
| 1 | 6 | 0.58 | 0.43 | 0.56 | 0.54 | 0.59 | 0.43 | 0.53 | 0.59651 | 0.56 | 0.53 | 0.56 | 0.56 | 0.53 | 0.58083 | 0.53 | 0.52 | 0.59 | 0.45 | 0.6 | 0.6 | 0.53 | 0.43
| Mean |  | 0.57833 | 0.475 | 0.58 | 0.54167 | 0.56333 | 0.435 | 0.525 | 0.58429 | 0.535 | 0.53833 | 0.54167 | 0.555 | 0.52333 | 0.5891 | 0.53667 | 0.54667 | 0.58833 | 0.46967 | 0.55 | 0.56667 | 0.50667 | 0.44167
| Median |  | 0.58 | 0.48 | 0.58 | 0.54 | 0.565 | 0.43 | 0.525 | 0.59475 | 0.53 | 0.535 | 0.54 | 0.56 | 0.525 | 0.58697 | 0.535 | 0.545 | 0.59 | 0.427 | 0.55 | 0.6 | 0.5 | 0.44
| Std Dev. |  | 0.0075277 | 0.02429 | 0.018974 | 0.0075277 | 0.043205 | 0.020736 | 0.013784 | 0.024643 | 0.016432 | 0.0098319 | 0.0098319 | 0.0083666 | 0.008165 | 0.011466 | 0.008165 | 0.019664 | 0.0075277 | 0.11716 | 0.054772 | 0.05164 | 0.012111 | 0.01169
| Rel.Std.Dev. |  | 1.302% | 5.114% | 3.271% | 1.39% | 7.67% | 4.767% | 2.626% | 4.218% | 3.071% | 1.826% | 1.815% | 1.507% | 1.56% | 1.946% | 1.521% | 3.597% | 1.28% | 24.94% | 9.959% | 9.113% | 2.39% | 2.647%
| PDM3 |  | 7.792% | -11.47% | 8.102% | 0.9576% | 4.996% | -18.92% | -2.149% | 8.901% | -0.285% | 0.3363% | 0.9576% | 3.443% | -2.459% | 9.797% | 0.02567% | 1.89% | 9.655% | -12.46% | 2.511% | 5.617% | -5.566% | -17.68%
| Z-Score (Absolute) |  | 0.91 | 1.55 | 0.95 | 0.04 | 0.56 | 2.5 | 0.36 | 1.05 | 0.12 | 0.04 | 0.04 | 0.36 | 0.4 | 1.17 | 0.08 | 0.16 | 1.15 | 1.67 | NA | NA | 0.79 | 2.34
Results from laboratories 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 550b (Certified Value 0.33 ppm) | [24 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 0.32 | 0.3 | 0.4 | 0.44928 | 0.3 | 0.32863 | 0.31 | 0.21 | [15 empty] | 1
| 1 | 2 | 0.33 | 0.3 | 0.4 | 0.43424 | 0.32 | 0.33646 | 0.33 | 0.23 | [15 empty] | 12
| 1 | 3 | 0.34 | 0.3 | 0.4 | 0.44072 | 0.35 | 0.33427 | 0.33 | 0.23 | [15 empty] | 16
| 1 | 4 | 0.33 | 0.3 | 0.4 | 0.45464 | 0.31 | 0.34272 | 0.31 | 0.23 | [15 empty] | 0.33177
| 1 | 5 | 0.31 | 0.3 | 0.3 | 0.4444 | 0.35 | 0.33055 | 0.32 | 0.23 | [15 empty] | 58
| 1 | 6 | 0.34 | 0.3 | 0.4 | 0.43464 | 0.33 | 0.33118 | 0.31 | 0.22
| Mean |  | 0.32833 | 0.3 | 0.38333 | 0.44299 | 0.32667 | 0.33397 | 0.31833 | 0.225
| Median |  | 0.33 | 0.3 | 0.4 | 0.44256 | 0.325 | 0.33272 | 0.315 | 0.23
| Std Dev. |  | 0.01169 | 0 | 0.040825 | 0.0081082 | 0.020656 | 0.0051152 | 0.0098319 | 0.0083666
| Rel.Std.Dev. |  | 3.561% | 0% | 10.65% | 1.83% | 6.323% | 1.532% | 3.089% | 3.718%
| PDM3 |  | -1.037% | -9.577% | 15.54% | 33.52% | -1.539% | 0.6618% | -4.051% | -32.18%
| Z-Score (Absolute) |  | 0.03 | 1.01 | 2.05 | 4.25 | 0.03 | 0.24 | 0.34 | 3.77
| Analytical results for U in OREAS 550b (Certified Value 8.83 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 8.82 | 8.1 | 9.8 | 8.38 | 11.4016 | 8.54 | 7.8 | 8.6 | 9.4705 | 8.6 | 9.1 | 8.9 | 8.9 | 8.7 | 9.4985 | 8.78 | 9.36 | 9.69 | 8.28 | 8.5 | 9.1 | 8.69 | 12 | 1
| 1 | 2 | 8.74 | 8.8 | 9.3 | 8.42 | 11.3872 | 8.06 | 8 | 8.9 | 9.6931 | 8.7 | 9.2 | 8.6 | 8.1 | 8.7 | 9.5144 | 8.91 | 9.67 | 9.33 | 8 | 8.4 | 9.2 | 9.01 | < 10 | 13
| 1 | 3 | 8.89 | 8.2 | 9.8 | 8.33 | 11.4248 | 8.71 | 8 | 8.3 | 9.2621 | 8.8 | 8.8 | 8.7 | 8.9 | 8.8 | 9.3088 | 8.93 | 9.21 | 9.6 | 7.83 | 8.5 | 9 | 8.8 | < 10 | 16
| 1 | 4 | 8.84 | 8.1 | 9.6 | 8.54 | 11.3848 | 8 | 8.3 | 8.6 | 9.6707 | 8.6 | 8.8 | 8.8 | 8.4 | 8.8 | 9.5479 | 8.76 | 8.72 | 9.58 | 9.64 | 8.3 | 9.2 | 8.86 | < 10 | 8.8288
| 1 | 5 | 8.67 | 7.9 | 9.1 | 8.52 | 11.3952 | 8.95 | 8.4 | 8.8 | 9.3773 | 8.3 | 9 | 8.6 | 9 | 8.8 | 9.3265 | 8.98 | 9.14 | 9.6 | 8.44 | 8.6 | 9.2 | 9 | < 10 | 59
| 1 | 6 | 8.91 | 7.4 | 9.6 | 8.35 | 11.444 | 8.54 | 8.1 | 8.4 | 9.2504 | 9 | 7 | 8.6 | 8.9 | 8.7 | 9.3869 | 8.99 | 9.04 | 9.61 | 8.74 | 8.7 | 9.3 | 9.04 | < 10
| Mean |  | 8.8117 | 8.0833 | 9.5333 | 8.4233 | 11.4063 | 8.4667 | 8.1 | 8.6 | 9.454 | 8.6667 | 8.65 | 8.7 | 8.7 | 8.75 | 9.4305 | 8.8917 | 9.19 | 9.5683 | 8.4883 | 8.5 | 9.1667 | 8.9 | 12
| Median |  | 8.83 | 8.1 | 9.6 | 8.4 | 11.3984 | 8.54 | 8.05 | 8.6 | 9.4239 | 8.65 | 8.9 | 8.65 | 8.9 | 8.75 | 9.4427 | 8.92 | 9.175 | 9.6 | 8.36 | 8.5 | 9.2 | 8.93 | 12
| Std Dev. |  | 0.091524 | 0.4535 | 0.28048 | 0.088242 | 0.023392 | 0.3706 | 0.21909 | 0.22804 | 0.19419 | 0.23381 | 0.82401 | 0.12649 | 0.36332 | 0.054772 | 0.10293 | 0.099079 | 0.31799 | 0.12287 | 0.6492 | 0.14142 | 0.10328 | 0.13957
| Rel.Std.Dev. |  | 1.039% | 5.61% | 2.942% | 1.048% | 0.2051% | 4.377% | 2.705% | 2.652% | 2.054% | 2.698% | 9.526% | 1.454% | 4.176% | 0.626% | 1.091% | 1.114% | 3.46% | 1.284% | 7.648% | 1.664% | 1.127% | 1.568%
| PDM3 |  | -0.1939% | -8.443% | 7.98% | -4.592% | 29.19% | -4.102% | -8.255% | -2.591% | 7.082% | -1.836% | -2.025% | -1.459% | -1.459% | -0.8924% | 6.815% | 0.7122% | 4.091% | 8.377% | -3.856% | -3.724% | 3.827% | 0.8066% | 35.92%
| Z-Score (Absolute) |  | 0.32 | 1.78 | 2.41 | 0.8 | 7.82 | 0.67 | 1.73 | 0.29 | 2.18 | 0.1 | 0.14 | 0 | 0 | 0.14 | 2.11 | 0.55 | 1.42 | 2.51 | 0.61 | 0.58 | 1.35 | 0.58 | 7.32
| Analytical results for V in OREAS 550b (Certified Value 201 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 205 | 199 | 190 | 197 | 197.48 | 187 | 196 | 194 | 213.4655 | 201 | 203 | 202 | 190 | 196 | 196.4265 | 204 | 194 | 200 | 190 | 206 | 201 | 206 | 217 | 1
| 1 | 2 | 204 | 201 | 188 | 194 | 197.35 | 202 | 196 | 184 | 211.6932 | 202 | 198 | 204 | 196 | 203 | 211.3032 | 206 | 198 | 210 | 202 | 204 | 201 | 202 | 217 | 28
| 1 | 3 | 210 | 202 | 193 | 195 | 204.11 | 214 | 196 | 195 | 209.9959 | 202 | 194 | 193 | 212 | 201 | 204.4162 | 208 | 201 | 200 | 203 | 195 | 201 | 200 | 214 | 16
| 1 | 4 | 208 | 199 | 191 | 197 | 203.55 | 208 | 198 | 187 | 208.8586 | 204 | 201 | 191 | 205 | 203 | 208.8812 | 206 | 199 | 209 | 200 | 199 | 200 | 201 | 217 | 200.6645
| 1 | 5 | 204 | 202 | 188 | 196 | 196.27 | 225 | 195 | 198 | 210.3145 | 198 | 199 | 195 | 203 | 202 | 207.947 | 208 | 200 | 199 | 199 | 207 | 201 | 199 | 219 | 60
| 1 | 6 | 205 | 196 | 199 | 196 | 200.18 | 229 | 196 | 185 | 213.008 | 204 | 194 | 198 | 208 | 195 | 210.9654 | 209 | 200 | 209 | 198 | 208 | 209 | 197 | 221
| Mean |  | 206 | 199.8333 | 191.5 | 195.8333 | 199.8233 | 210.8333 | 196.1667 | 190.5 | 211.2226 | 201.8333 | 198.1667 | 197.1667 | 202.3333 | 200 | 206.6566 | 206.8333 | 198.6667 | 204.5 | 198.6667 | 203.1667 | 202.1667 | 200.8333 | 217.5
| Median |  | 205 | 200 | 190.5 | 196 | 198.83 | 211 | 196 | 190.5 | 211.0039 | 202 | 198.5 | 196.5 | 204 | 201.5 | 208.4141 | 207 | 199.5 | 204.5 | 199.5 | 205 | 201 | 200.5 | 217
| Std Dev. |  | 2.4495 | 2.3166 | 4.1352 | 1.169 | 3.3653 | 15.4585 | 0.98319 | 5.8907 | 1.8087 | 2.2286 | 3.656 | 5.1153 | 8.0664 | 3.5777 | 5.5936 | 1.8348 | 2.5033 | 5.3198 | 4.6332 | 5.1153 | 3.3714 | 3.0605 | 2.3452
| Rel.Std.Dev. |  | 1.189% | 1.159% | 2.159% | 0.597% | 1.684% | 7.332% | 0.5012% | 3.092% | 0.8563% | 1.104% | 1.845% | 2.594% | 3.987% | 1.789% | 2.707% | 0.8871% | 1.26% | 2.601% | 2.332% | 2.518% | 1.668% | 1.524% | 1.078%
| PDM3 |  | 2.659% | -0.4142% | -4.567% | -2.408% | -0.4192% | 5.068% | -2.241% | -5.065% | 5.262% | 0.5825% | -1.245% | -1.743% | 0.8316% | -0.3312% | 2.986% | 3.074% | -0.9956% | 1.911% | -0.9956% | 1.247% | 0.7486% | 0.08413% | 8.39%
| Z-Score (Absolute) |  | 0.95 | 0.18 | 1.72 | 0.92 | 0.19 | 1.84 | 0.86 | 1.9 | 1.91 | 0.18 | 0.49 | 0.67 | 0.28 | 0.15 | 1.07 | 1.1 | 0.4 | 0.67 | 0.4 | 0.43 | 0.25 | 0 | 3.07
| Analytical results for W in OREAS 550b (Certified Value 2.71 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
2
| 1 | 1 | 2.8 | 2.6 | 3.6 | 2.3 | 2.6 | 3.4 | 2.3 | 2.459 | 4.2 | 2.8 | 3.8 | 2.6 | 2.9 | 3.0191 | 2.7 | 2.7 | 2.7 | 3.32 | 3.1 | 2.6 | 3 | 2.5 |  | 1
| 1 | 2 | 2.8 | 2.7 | 3.8 | 2.2 | 2.4 | 3.4 | 2.4 | 2.1004 | 3 | 2.8 | 3.1 | 4 | 2.7 | 3.2277 | 2.75 | 2.7 | 2.8 | 5.7 | 2.4 | 2.4 | 3.1 | 2.5 |  | 14
| 1 | 3 | 2.6 | 3.6 | 3.8 | 2.2 | 2.6 | 2.9 | 2.1 | 2.0243 | 2.9 | 3 | 3 | 3 | 3.2 | 2.9721 | 2.8 | 2.6 | 2.5 | 5.09 | 2.5 | 3.1 | 2.9 | 2.6 |  | 16
| 1 | 4 | 2.6 | 2.6 | 4.1 | 2.1 | 2.5 | 3.3 | 2.2 | 2.5849 | 3 | 3.1 | 2.8 | 2.8 | 2.6 | 3.1322 | 2.72 | 2.6 | 2.7 | 4.1 | 2.3 | 2.4 | 3 | 2.6 |  | 2.7063
| 1 | 5 | 2.8 | 2.5 | 5.4 | 2.2 | 2.5 | 3.4 | 2.3 | 1.6298 | 3.1 | 2.7 | 2.9 | 2.6 | 2.9 | 3.1447 | 2.95 | 2.3 | 2.4 | 3.1 | 2.4 | 2.3 | 3.3 | 3.2 |  | 61
| 1 | 6 | 2.7 | 2.4 | 3.9 | 2.2 | 2.4 | 3.1 | 2.2 | 1.9275 | 2.9 | 2.9 | 2.9 | 2.6 | 2.7 | 2.9725 | 2.85 | 2.6 | 2.7 | 5.1 | 2.4 | 3 | 3.5 | 2.8
| Mean |  | 2.7167 | 2.7333 | 4.1 | 2.2 | 2.5 | 3.25 | 2.25 | 2.121 | 3.1833 | 2.8833 | 3.0833 | 2.9333 | 2.8333 | 3.078 | 2.795 | 2.5833 | 2.6333 | 4.4017 | 2.5167 | 2.6333 | 3.1333 | 2.7
| Median |  | 2.75 | 2.6 | 3.85 | 2.2 | 2.5 | 3.35 | 2.25 | 2.0624 | 3 | 2.85 | 2.95 | 2.7 | 2.8 | 3.0756 | 2.775 | 2.6 | 2.7 | 4.595 | 2.4 | 2.5 | 3.05 | 2.6
| Std Dev. |  | 0.098319 | 0.43665 | 0.65727 | 0.063246 | 0.089443 | 0.20736 | 0.10488 | 0.35156 | 0.50365 | 0.1472 | 0.3656 | 0.5465 | 0.21602 | 0.10545 | 0.093541 | 0.1472 | 0.15055 | 1.0585 | 0.29269 | 0.33862 | 0.22509 | 0.26833
| Rel.Std.Dev. |  | 3.619% | 15.98% | 16.03% | 2.875% | 3.578% | 6.38% | 4.661% | 16.58% | 15.82% | 5.105% | 11.86% | 18.63% | 7.624% | 3.426% | 3.347% | 5.698% | 5.717% | 24.05% | 11.63% | 12.86% | 7.184% | 9.938%
| PDM3 |  | 0.3838% | 0.9996% | 51.5% | -18.71% | -7.622% | 20.09% | -16.86% | -21.63% | 17.63% | 6.542% | 13.93% | 8.39% | 4.695% | 13.74% | 3.278% | -4.543% | -2.695% | 62.65% | -7.006% | -2.695% | 15.78% | -0.2321%
| Z-Score (Absolute) |  | 0.13 | 0.08 | 3.52 | 1.49 | 0.7 | 1.28 | 1.36 | 1.7 | 1.1 | 0.31 | 0.84 | 0.45 | 0.18 | 0.83 | 0.08 | 0.48 | 0.34 | 4.32 | 0.65 | 0.34 | 0.97 | 0.17
| Analytical results for Y in OREAS 550b (Certified Value 19.1 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 19.48 | 18.6 | 21.7 | 18.9 | 26.1418 | 16 | 18.5 | 18.8 | 19.5909 | 18.5 | 19.3 | 19.3 | 21.1 | 19.4 | 19.6041 | 19 | 19.2 | 19.89 | 17.78 | 19 | 19.6 | 20.2 | 15.8 | 1
| 1 | 2 | 19.33 | 18.7 | 21.6 | 18.7 | 26.1558 | 16.5 | 19 | 19.8 | 20.4992 | 19.6 | 19.3 | 18.8 | 19.5 | 19.7 | 19.8208 | 19 | 19 | 20.12 | 18.9 | 18.3 | 20.1 | 19.5 | 17.4 | 15
| 1 | 3 | 18.68 | 18.8 | 21.6 | 18.5 | 26.145 | 16.9 | 18.6 | 18.9 | 19.1207 | 18.9 | 19.4 | 18.9 | 21.5 | 19.4 | 19.7788 | 19.2 | 18.3 | 19.66 | 19.02 | 17.7 | 19.7 | 20.1 | 17 | 16
| 1 | 4 | 18.86 | 18.8 | 21.2 | 19 | 26.2505 | 16.6 | 19.5 | 20.1 | 20.3456 | 18.6 | 19.9 | 18.9 | 20.8 | 19.9 | 19.9028 | 19.6 | 18.6 | 19.91 | 18.37 | 18 | 20.5 | 18.9 | 17 | 19.0781
| 1 | 5 | 18.45 | 18.7 | 21 | 18.7 | 26.1723 | 17.5 | 19.3 | 18.8 | 21.2261 | 18.1 | 19.2 | 19 | 21 | 20 | 19.8767 | 19.8 | 18.8 | 20.62 | 18.4 | 19.1 | 20.3 | 18.8 | 16.8 | 62
| 1 | 6 | 19.82 | 18.8 | 22 | 18.8 | 26.2194 | 17 | 19.4 | 20.3 | 20.2183 | 19.2 | 17.4 | 18.5 | 21 | 20.1 | 20.1756 | 19.3 | 18.2 | 20.23 | 18.3 | 18.5 | 22.1 | 18.7 | 16.9
| Mean |  | 19.1033 | 18.7333 | 21.5167 | 18.7667 | 26.1808 | 16.75 | 19.05 | 19.45 | 20.1668 | 18.8167 | 19.0833 | 18.9 | 20.8167 | 19.75 | 19.8598 | 19.3167 | 18.6833 | 20.0717 | 18.4617 | 18.4333 | 20.3833 | 19.3667 | 16.8167
| Median |  | 19.095 | 18.75 | 21.6 | 18.75 | 26.1641 | 16.75 | 19.15 | 19.35 | 20.282 | 18.75 | 19.3 | 18.9 | 21 | 19.8 | 19.8487 | 19.25 | 18.7 | 20.015 | 18.385 | 18.4 | 20.2 | 19.2 | 16.95
| Std Dev. |  | 0.52386 | 0.08165 | 0.36009 | 0.17512 | 0.044402 | 0.50892 | 0.42308 | 0.69498 | 0.73394 | 0.53448 | 0.8612 | 0.26077 | 0.68532 | 0.30166 | 0.18726 | 0.32506 | 0.392 | 0.33367 | 0.44875 | 0.55015 | 0.90866 | 0.66833 | 0.53821
| Rel.Std.Dev. |  | 2.742% | 0.4359% | 1.674% | 0.9331% | 0.1696% | 3.038% | 2.221% | 3.573% | 3.639% | 2.84% | 4.513% | 1.38% | 3.292% | 1.527% | 0.9429% | 1.683% | 2.098% | 1.662% | 2.431% | 2.985% | 4.458% | 3.451% | 3.2%
| PDM3 |  | 0.1323% | -1.807% | 12.78% | -1.632% | 37.23% | -12.2% | -0.1472% | 1.949% | 5.707% | -1.37% | 0.02747% | -0.9335% | 9.113% | 3.522% | 4.097% | 1.251% | -2.069% | 5.208% | -3.231% | -3.38% | 6.842% | 1.513% | -11.85%
| Z-Score (Absolute) |  | 0 | 0.39 | 2.54 | 0.35 | 7.44 | 2.47 | 0.06 | 0.36 | 1.12 | 0.3 | 0.02 | 0.21 | 1.8 | 0.68 | 0.79 | 0.22 | 0.44 | 1.02 | 0.67 | 0.7 | 1.35 | 0.28 | 2.4
| Analytical results for Yb in OREAS 550b (Certified Value 2.22 ppm) | [24 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 15 | 16 | 17 | 25 | 26 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 2.28 | 2 | 2.5 | 2.8149 | 2.1 | 2.1 | 2.319 | 2.15 | 2.2 | 2.11 | 1.6 | [12 empty] | 1
| 1 | 2 | 2.25 | 2.2 | 2.5 | 2.8128 | 2.2 | 2.3 | 2.3738 | 2.23 | 2.1 | 2.24 | 1.7 | [12 empty] | 16
| 1 | 3 | 2.17 | 2 | 2.5 | 2.8045 | 2.4 | 2 | 2.2966 | 2.21 | 2.2 | 2.11 | 1.8 | [12 empty] | 16
| 1 | 4 | 2.22 | 2 | 2.5 | 2.8148 | 2.2 | 2.2 | 2.4308 | 2.24 | 2.1 | 2.1 | 1.7 | [12 empty] | 2.2163
| 1 | 5 | 2.22 | 2 | 2.4 | 2.8064 | 2.4 | 1.9 | 2.2844 | 2.18 | 2.2 | 2.17 | 1.7 | [12 empty] | 63
| 1 | 6 | 2.23 | 1.8 | 2.5 | 2.8149 | 2.3 | 2.3 | 2.2656 | 2.26 | 2.1 | 2.14 | 1.7
| Mean |  | 2.2283 | 2 | 2.4833 | 2.8114 | 2.2667 | 2.1333 | 2.3284 | 2.2117 | 2.15 | 2.145 | 1.7
| Median |  | 2.225 | 2 | 2.5 | 2.8138 | 2.25 | 2.15 | 2.3078 | 2.22 | 2.15 | 2.125 | 1.7
| Std Dev. |  | 0.03656 | 0.12649 | 0.040825 | 0.0047037 | 0.12111 | 0.1633 | 0.062513 | 0.040702 | 0.054772 | 0.053198 | 0.063246
| Rel.Std.Dev. |  | 1.641% | 6.325% | 1.644% | 0.1673% | 5.343% | 7.655% | 2.685% | 1.84% | 2.548% | 2.48% | 3.72%
| PDM3 |  | 0.5429% | -9.76% | 12.05% | 26.85% | 2.273% | -3.743% | 5.056% | -0.2091% | -2.991% | -3.217% | -23.3%
| Z-Score (Absolute) |  | 0.14 | 1.82 | 2.34 | 5.16 | 0.47 | 0.67 | 1 | 0 | 0.53 | 0.57 | 4.4
| Analytical results for Zn in OREAS 550b (Certified Value 26.6 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 30 | 43.5 | 28 | 27 | 34.95 | 21 | 29 | 25 | 26.3048 | 26 | 30 | 26 | 26 | 26 | 26.4275 | 27.5 | 26 | 29 | 39 | 25 | 27 | 26.1 | 26 | 1
| 1 | 2 | 27 | 47.8 | 28 | 27 | 30.62 | 23 | 30 | 24 | 27.4625 | 26 | 30 | 25 | 25 | 27 | 26.3919 | 28.4 | 27 | 30 | 26 | 25 | 27 | 25.9 | 27 | 17
| 1 | 3 | 27 | 42.5 | 28 | 27 | 32.11 | 23 | 33 | 23 | 26.3721 | 25 | 30 | 23 | 27 | 27 | 26.3479 | 26.9 | 26 | 29 | 25 | 27 | 27 | 24.5 | 27 | 16
| 1 | 4 | 26 | 43.3 | 26 | 27 | 30.72 | 24 | 29 | 24 | 27.8452 | 25 | 30 | 23 | 27 | 26 | 26.3117 | 27.2 | 27 | 29 | 25 | 25 | 27 | 24.7 | 25 | 26.6105
| 1 | 5 | 26 | 41.8 | 29 | 27 | 32.32 | 27 | 32 | 25 | 27.4549 | 24 | 30 | 24 | 27 | 26 | 26.1039 | 27.7 | 27 | 29 | 25 | 26 | 28 | 26.7 | 27 | 64
| 1 | 6 | 29 | 38.7 | 27 | 27 | 33.82 | 16 | 30 | 23 | 26.3867 | 25 | 29 | 24 | 28 | 27 | 26.7833 | 27 | 26 | 29 | 25 | 25 | 29 | 26.6 | 28
| Mean |  | 27.5 | 42.9333 | 27.6667 | 27 | 32.4233 | 22.3333 | 30.5 | 24 | 26.971 | 25.1667 | 29.8333 | 24.1667 | 26.6667 | 26.5 | 26.3944 | 27.45 | 26.5 | 29.1667 | 27.5 | 25.5 | 27.5 | 25.75 | 26.6667
| Median |  | 27 | 42.9 | 28 | 27 | 32.215 | 23 | 30 | 24 | 26.9208 | 25 | 30 | 24 | 27 | 26.5 | 26.3699 | 27.35 | 26.5 | 29 | 25 | 25 | 27 | 26 | 27
| Std Dev. |  | 1.6432 | 2.9507 | 1.0328 | 0 | 1.7086 | 3.6697 | 1.6432 | 0.89443 | 0.6905 | 0.75277 | 0.40825 | 1.169 | 1.0328 | 0.54772 | 0.22171 | 0.55408 | 0.54772 | 0.40825 | 5.648 | 0.83666 | 0.83666 | 0.94181 | 1.0328
| Rel.Std.Dev. |  | 5.975% | 6.873% | 3.733% | 0% | 5.27% | 16.43% | 5.387% | 3.727% | 2.56% | 2.991% | 1.368% | 4.837% | 3.873% | 2.067% | 0.84% | 2.018% | 2.067% | 1.4% | 20.54% | 3.281% | 3.042% | 3.658% | 3.873%
| PDM3 |  | 3.343% | 61.34% | 3.969% | 1.464% | 21.84% | -16.07% | 14.62% | -9.81% | 1.355% | -5.426% | 12.11% | -9.184% | 0.2112% | -0.4151% | -0.8121% | 3.155% | -0.4151% | 9.606% | 3.343% | -4.173% | 3.343% | -3.234% | 0.2112%
| Z-Score (Absolute) |  | 0.51 | 15.47 | 0.67 | 0.03 | 5.29 | 4.5 | 3.42 | 2.88 | 0 | 1.75 | 2.77 | 2.72 | 0.3 | 0.46 | 0.56 | 0.46 | 0.46 | 2.13 | 0.51 | 1.43 | 0.51 | 1.18 | 0.3
| Analytical results for Zr in OREAS 550b (Certified Value 134 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 131.3 | 123.6 | 140 | 136 | 122.33 | 122 | 131 | 143.7 | 129.1233 | 138.5 | 138 | 129.5 | 125.5 | 132.5 | 135.3455 | 134 | 133 | 136.4 | 122.1 | 128.2 | 127.3 | 143 | 138 | 1
| 1 | 2 | 131.9 | 124.1 | 139 | 135 | 120.37 | 128 | 136 | 148.9 | 131.2236 | 142 | 138 | 152 | 126 | 139 | 132.8227 | 136 | 135 | 135.3 | 128.9 | 126.9 | 127.4 | 139 | 141 | 22
| 1 | 3 | 130.8 | 125.4 | 142 | 135 | 121.32 | 133 | 132 | 145.2 | 136.1694 | 142.5 | 140.5 | 127.5 | 137 | 138 | 133.2661 | 137 | 135 | 136.8 | 133 | 129.1 | 127.8 | 139 | 140 | 16
| 1 | 4 | 130.3 | 123.8 | 138 | 135 | 121.48 | 131 | 140 | 145.3 | 131.634 | 140 | 141.5 | 128 | 135 | 137 | 137.04 | 136 | 133 | 136.3 | 125.9 | 125.3 | 129.2 | 134 | 137 | 134.4601
| 1 | 5 | 130.3 | 125 | 137 | 135 | 120.5 | 142 | 139 | 148.7 | 132.7786 | 141.5 | 140 | 128.5 | 134 | 139 | 136.3566 | 138 | 135 | 131.5 | 125.5 | 126.3 | 127.9 | 136 | 140 | 65
| 1 | 6 | 133.1 | 120.9 | 142 | 133 | 120.52 | 144 | 138 | 143.9 | 136.3584 | 143.5 | 138 | 127 | 135 | 138 | 139.0117 | 136 | 135 | 140.1 | 127.2 | 130.7 | 128.7 | 136 | 140
| Mean |  | 131.2833 | 123.8 | 139.6667 | 134.8333 | 121.0867 | 133.3333 | 136 | 145.95 | 132.8812 | 141.3333 | 139.3333 | 132.0833 | 132.0833 | 137.25 | 135.6404 | 136.1667 | 134.3333 | 136.0667 | 127.1 | 127.75 | 128.05 | 137.8333 | 139.3333
| Median |  | 131.05 | 123.95 | 139.5 | 135 | 120.92 | 132 | 137 | 145.25 | 132.2063 | 141.75 | 139 | 128.25 | 134.5 | 138 | 135.8511 | 136 | 135 | 136.35 | 126.55 | 127.55 | 127.85 | 137.5 | 140
| Std Dev. |  | 1.0815 | 1.5837 | 2.0656 | 0.98319 | 0.76602 | 8.3825 | 3.7417 | 2.3028 | 2.8754 | 1.8074 | 1.5384 | 9.795 | 5.0042 | 2.4444 | 2.3453 | 1.3292 | 1.0328 | 2.7718 | 3.6622 | 1.9776 | 0.75033 | 3.1885 | 1.5055
| Rel.Std.Dev. |  | 0.8238% | 1.279% | 1.479% | 0.7292% | 0.6326% | 6.287% | 2.751% | 1.578% | 2.164% | 1.279% | 1.104% | 7.416% | 3.789% | 1.781% | 1.729% | 0.9761% | 0.7688% | 2.037% | 2.881% | 1.548% | 0.586% | 2.313% | 1.081%
| PDM3 |  | -2.363% | -7.928% | 3.872% | 0.2776% | -9.946% | -0.838% | 1.145% | 8.545% | -1.174% | 5.112% | 3.624% | -1.768% | -1.768% | 2.075% | 0.8779% | 1.269% | -0.09426% | 1.195% | -5.474% | -4.99% | -4.767% | 2.509% | 3.624%
| Z-Score (Absolute) |  | 0.8 | 2.48 | 1.09 | 0 | 3.09 | 0.34 | 0.26 | 2.5 | 0.44 | 1.46 | 1.01 | 0.62 | 0.62 | 0.54 | 0.18 | 0.3 | 0.11 | 0.28 | 1.74 | 1.59 | 1.52 | 0.67 | 1.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C813: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C887: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C905: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C943: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C961: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1124: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1142: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach (05%)
| Analytical results for Co in OREAS 550b (Certified Value 0.042 wt.%) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 19 | 23 | 25 | [11 empty] | wt.%
|  |  | AL-5S*OES | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | [11 empty] | 3
3
| 1 | 1 | 0.0388 | 0.0447 | 0.048666 | 0.041406 | 0.0388 | 0.045538 | 0.043 | 0.0358 | 0.6371 | [11 empty] | 1
| 1 | 2 | 0.0397 | 0.0445 | 0.048566 | 0.040735 | 0.0386 | 0.046956 | 0.043 | 0.0361 | 0.6172 | [11 empty] | 29
| 1 | 3 | 0.0385 | 0.0442 | 0.048582 | 0.041254 | 0.0388 | 0.045927 | 0.042 | 0.0355 | 0.6636 | [11 empty] | 16
| 1 | 4 | 0.039 | 0.0444 | 0.048805 | 0.041395 | 0.0388 | 0.045973 | 0.042 | 0.0352 | 0.6512 | [11 empty] | 0.042001
| 1 | 5 | 0.0381 | 0.0444 | 0.048022 | 0.041198 | 0.0384 | 0.046865 | 0.041 | 0.0366 | 0.6459 | [11 empty] | 67
| 1 | 6 | 0.0389 | 0.0445 | 0.048586 | 0.042092 | 0.0389 | 0.044978 | 0.042 | 0.0363 | 0.6457
| Mean |  | 0.038833 | 0.04445 | 0.048538 | 0.041347 | 0.038717 | 0.04604 | 0.042167 | 0.035917 | 0.64345
| Median |  | 0.03885 | 0.04445 | 0.048584 | 0.041325 | 0.0388 | 0.04595 | 0.042 | 0.03595 | 0.6458
| Std Dev. |  | 0.00053541 | 0.00016432 | 0.00026803 | 0.00043978 | 0.00018348 | 0.00076359 | 0.00075277 | 0.00051929 | 0.015534
| Rel.Std.Dev. |  | 1.379% | 0.3697% | 0.5522% | 1.064% | 0.4739% | 1.659% | 1.785% | 1.446% | 2.414%
| PDM3 |  | -7.542% | 5.831% | 15.56% | -1.558% | -7.819% | 9.616% | 0.3946% | -14.49% | 1432%
| Z-Score (Absolute) |  | 0.65 | 0.45 | 1.25 | 0.16 | 0.67 | 0.76 | 0 | 1.22 | 117.52
| Analytical results for Cu in OREAS 550b (Certified Value 0.832 wt.%) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 19 | 20 | 23 | 25 | 26 | wt.%
|  |  | AL-5S*OES | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | 3
3
| 1 | 1 | 0.8369 | 0.794 | 0.889 | 0.856 | 1.1959 | 0.81376 | 0.9299 | 0.868 | 0.80588 | 0.745 | 0.803 | 0.796 | 0.803 | 0.77367 | 0.862 | 0.845 | 0.831 | 1.01 | 8.613 | 0.848 | 1
| 1 | 2 | 0.8559 | 0.809 | 0.88 | 0.85 | 1.1793 | 0.80639 | 0.9241 | 0.865 | 0.81455 | 0.764 | 0.806 | 0.788 | 0.814 | 0.77052 | 0.863 | 0.828 | 0.83 | 1.04 | 7.941 | 0.821 | 30
| 1 | 3 | 0.8331 | 0.822 | 0.881 | 0.863 | 1.1823 | 0.83734 | 0.9174 | 0.864 | 0.8042 | 0.771 | 0.805 | 0.786 | 0.944 | 0.77166 | 0.862 | 0.798 | 0.826 | 1.03 | 8.193 | 0.845 | 16
| 1 | 4 | 0.8256 | 0.804 | 0.89 | 0.864 | 1.2042 | 0.81374 | 0.938 | 0.866 | 0.8252 | 0.767 | 0.793 | 0.738 | 0.897 | 0.7759 | 0.871 | 0.822 | 0.823 | 1.02 | 8.038 | 0.841 | 0.83217
| 1 | 5 | 0.8213 | 0.811 | 0.884 | 0.857 | 1.1802 | 0.83414 | 0.9259 | 0.868 | 0.82365 | 0.78 | 0.796 | 0.775 | 0.91 | 0.77636 | 0.861 | 0.802 | 0.819 | 1.05 | 8.276 | 0.844 | 68
| 1 | 6 | 0.8351 | 0.794 | 0.887 | 0.851 | 1.1933 | 0.8125 | 0.9258 | 0.863 | 0.81089 | 0.776 | 0.802 | 0.784 | 0.843 | 0.77267 | 0.865 | 0.817 | 0.835 | 1.04 | 8.212 | 0.844
| Mean |  | 0.83465 | 0.80567 | 0.88517 | 0.85683 | 1.1892 | 0.81964 | 0.92685 | 0.86567 | 0.81406 | 0.76717 | 0.80083 | 0.77783 | 0.8685 | 0.77346 | 0.864 | 0.81867 | 0.82733 | 1.0317 | 8.2122 | 0.8405
| Median |  | 0.8341 | 0.8065 | 0.8855 | 0.8565 | 1.1878 | 0.81375 | 0.92585 | 0.8655 | 0.81272 | 0.769 | 0.8025 | 0.785 | 0.87 | 0.77317 | 0.8625 | 0.8195 | 0.828 | 1.035 | 8.2025 | 0.844
| Std Dev. |  | 0.011995 | 0.010783 | 0.0041673 | 0.0058452 | 0.010136 | 0.012803 | 0.006819 | 0.0020656 | 0.008839 | 0.012319 | 0.0051929 | 0.020653 | 0.056825 | 0.0023206 | 0.0036878 | 0.017317 | 0.0058195 | 0.01472 | 0.23194 | 0.0098133
| Rel.Std.Dev. |  | 1.437% | 1.338% | 0.4708% | 0.6822% | 0.8523% | 1.562% | 0.7357% | 0.2386% | 1.086% | 1.606% | 0.6484% | 2.655% | 6.543% | 0.3% | 0.4268% | 2.115% | 0.7034% | 1.427% | 2.824% | 1.168%
| PDM3 |  | 0.2984% | -3.184% | 6.369% | 2.964% | 42.91% | -1.505% | 11.38% | 4.026% | -2.176% | -7.811% | -3.765% | -6.529% | 4.366% | -7.054% | 3.825% | -1.622% | -0.5808% | 23.97% | 886.8% | 1.001%
| Z-Score (Absolute) |  | 0.06 | 0.68 | 1.02 | 0.41 | 7.55 | 0.38 | 1.92 | 0.6 | 0.5 | 1.51 | 0.79 | 1.28 | 0.66 | 1.38 | 0.57 | 0.41 | 0.22 | 4.17 | 158.28 | 0.06
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 550b (Certified Value 3.12 wt.%) | [24 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 3.13 | 3.05 | 3.17 | 3.14 | 3.1 | 3.13 | 3.06 | 2.8802 | 3.22 | 3.08 | 3.09 | 3.05 | 3.15 | 3.0164 | 3.256 | 2.83 | 3.25 | 3.14 | 2.97 | 2.89 | 3.2 | 3.13 | 3.21 | 1
| 1 | 2 | 3.13 | 3 | 3.05 | 3.14 | 3.18 | 3.14 | 3.05 | 2.9515 | 3.2 | 2.99 | 3.08 | 2.99 | 3.16 | 3.0495 | 3.31 | 2.86 | 3.3 | 3.13 | 2.94 | 2.87 | 3.14 | 3.14 | 3.16 | #N/A
| 1 | 3 | 3.15 | 3.04 | 2.99 | 3.18 | 3.17 | 3.15 | 3.07 | 2.9238 | 3.17 | 3.11 | 3.1 | 3.05 | 3.14 | 3.0235 | 3.277 | 2.81 | 3.23 | 3.14 | 3.03 | 2.87 | 3.25 | 3.13 | 3.11 | 16
| 1 | 4 | 3.1 | 3.02 | 3.11 | 3.18 | 3.05 | 3.15 | 3.05 | 2.9071 | 3.19 | 3.06 | 3.06 | 3.03 | 3.13 | 3.0385 | 3.288 | 2.84 | 3.22 | 3.13 | 3.02 | 2.87 | 3.18 | 3.15 | 3.17 | 3.1219
| 1 | 5 | 3.06 | 3.02 | 3.09 | 3.17 | 3.17 | 3.14 | 3.06 | 2.8779 | 3.23 | 3.01 | 3.06 | 3.04 | 3.14 | 3.0149 | 3.276 | 2.81 | 3.26 | 3.14 | 3 | 2.88 | 3.16 | 3.12 | 3.13 | 70
| 1 | 6 | 3.09 | 3.02 | 3.13 | 3.16 | 3.2 | 3.14 | 3.07 | 2.9633 | 3.21 | 3.13 | 3.1 | 3.06 | 3.13 | 3.034 | 3.323 | 2.8 | 3.29 | 3.13 | 3.02 | 2.89 | 3.18 | 3.15 | 3.11
| Mean |  | 3.11 | 3.025 | 3.09 | 3.1617 | 3.145 | 3.1417 | 3.06 | 2.9173 | 3.2033 | 3.0633 | 3.0817 | 3.0367 | 3.1417 | 3.0295 | 3.2883 | 2.825 | 3.2583 | 3.135 | 2.9967 | 2.8783 | 3.185 | 3.1367 | 3.1483
| Median |  | 3.115 | 3.02 | 3.1 | 3.165 | 3.17 | 3.14 | 3.06 | 2.9154 | 3.205 | 3.07 | 3.085 | 3.045 | 3.14 | 3.0287 | 3.2825 | 2.82 | 3.255 | 3.135 | 3.01 | 2.875 | 3.18 | 3.135 | 3.145
| Std Dev. |  | 0.032863 | 0.017607 | 0.063246 | 0.018348 | 0.057533 | 0.0075277 | 0.0089443 | 0.035674 | 0.021602 | 0.055015 | 0.018348 | 0.025033 | 0.01169 | 0.013598 | 0.024484 | 0.022583 | 0.031885 | 0.0054772 | 0.035024 | 0.0098319 | 0.037815 | 0.012111 | 0.0392
| Rel.Std.Dev. |  | 1.057% | 0.582% | 2.047% | 0.5803% | 1.829% | 0.2396% | 0.2923% | 1.223% | 0.6744% | 1.796% | 0.5954% | 0.8244% | 0.3721% | 0.4489% | 0.7446% | 0.7994% | 0.9786% | 0.1747% | 1.169% | 0.3416% | 1.187% | 0.3861% | 1.245%
| PDM3 |  | -0.3808% | -3.104% | -1.021% | 1.274% | 0.7403% | 0.6335% | -1.982% | -6.554% | 2.609% | -1.876% | -1.288% | -2.73% | 0.6335% | -2.961% | 5.332% | -9.51% | 4.371% | 0.42% | -4.011% | -7.802% | 2.022% | 0.4733% | 0.847%
| Z-Score (Absolute) |  | 0 | 1.11 | 0.26 | 0.67 | 0.46 | 0.41 | 0.65 | 2.52 | 1.22 | 0.61 | 0.37 | 0.96 | 0.41 | 1.05 | 2.33 | 3.72 | 1.94 | 0.33 | 1.48 | 3.02 | 0.98 | 0.35 | 0.5
| Analytical results for S in OREAS 550b (Certified Value 1.17 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 1.18 | 1.21 | 1.22 | 1.28 | 1.19 | 1.17 | 1.15 | 1.1668 | 1.19 | 0.96 | 1.23 | 1.16 | 1.2 | 1.1394 | 1.142 | 1.2 | 1.21 | 1.13 | 1.14 | 1.19 | 1.16 | 1.14 | 1.15 | 1
| 1 | 2 | 1.18 | 1.22 | 1.17 | 1.26 | 1.17 | 1.18 | 1.13 | 1.1526 | 1.18 | 1.17 | 1.2 | 1.12 | 1.2 | 1.1496 | 1.153 | 1.23 | 1.22 | 1.12 | 1 | 1.2 | 1.13 | 1.15 | 1.16 | 30
| 1 | 3 | 1.18 | 1.21 | 1.18 | 1.26 | 1.14 | 1.17 | 1.16 | 1.14 | 1.21 | 1.14 | 1.2 | 1.16 | 1.2 | 1.1471 | 1.137 | 1.19 | 1.23 | 1.13 | 1.06 | 1.21 | 1.19 | 1.14 | 1.14 | 16
| 1 | 4 | 1.18 | 1.19 | 1.2 | 1.28 | 1.06 | 1.17 | 1.15 | 1.1311 | 1.2 | 1.02 | 1.2 | 1.14 | 1.2 | 1.1436 | 1.15 | 1.17 | 1.25 | 1.13 | 1.02 | 1.21 | 1.16 | 1.14 | 1.14 | 1.1657
| 1 | 5 | 1.19 | 1.21 | 1.2 | 1.27 | 1.14 | 1.18 | 1.13 | 1.1681 | 1.16 | 1.15 | 1.18 | 1.16 | 1.19 | 1.1473 | 1.148 | 1.17 | 1.19 | 1.13 | 1.1 | 1.21 | 1.16 | 1.15 | 1.13 | 71
| 1 | 6 | 1.2 | 1.17 | 1.19 | 1.27 | 1.13 | 1.18 | 1.16 | 1.1428 | 1.19 | 1.08 | 1.19 | 1.15 | 1.19 | 1.1359 | 1.12 | 1.18 | 1.23 | 1.13 | 1.13 | 1.21 | 1.16 | 1.15 | 1.14
| Mean |  | 1.185 | 1.2017 | 1.1933 | 1.27 | 1.1383 | 1.175 | 1.1467 | 1.1502 | 1.1883 | 1.0867 | 1.2 | 1.1483 | 1.1967 | 1.1438 | 1.1417 | 1.19 | 1.2217 | 1.1283 | 1.075 | 1.205 | 1.16 | 1.145 | 1.1433
| Median |  | 1.18 | 1.21 | 1.195 | 1.27 | 1.14 | 1.175 | 1.15 | 1.1477 | 1.19 | 1.11 | 1.2 | 1.155 | 1.2 | 1.1454 | 1.145 | 1.185 | 1.225 | 1.13 | 1.08 | 1.21 | 1.16 | 1.145 | 1.14
| Std Dev. |  | 0.0083666 | 0.018348 | 0.017512 | 0.0089443 | 0.04446 | 0.0054772 | 0.013663 | 0.015016 | 0.017224 | 0.082865 | 0.016733 | 0.016021 | 0.005164 | 0.0052861 | 0.012078 | 0.022804 | 0.020412 | 0.0040825 | 0.057879 | 0.0083666 | 0.018974 | 0.0054772 | 0.010328
| Rel.Std.Dev. |  | 0.706% | 1.527% | 1.467% | 0.7043% | 3.906% | 0.4661% | 1.192% | 1.305% | 1.449% | 7.626% | 1.394% | 1.395% | 0.4315% | 0.4621% | 1.058% | 1.916% | 1.671% | 0.3618% | 5.384% | 0.6943% | 1.636% | 0.4784% | 0.9033%
| PDM3 |  | 1.658% | 3.087% | 2.373% | 8.95% | -2.346% | 0.7998% | -1.631% | -1.325% | 1.944% | -6.778% | 2.945% | -1.488% | 2.659% | -1.875% | -2.06% | 2.087% | 4.803% | -3.204% | -7.779% | 3.373% | -0.487% | -1.774% | -1.917%
| Z-Score (Absolute) |  | 0.59 | 0.99 | 0.79 | 2.62 | 0.52 | 0.36 | 0.32 | 0.23 | 0.67 | 1.75 | 0.95 | 0.28 | 0.87 | 0.39 | 0.44 | 0.71 | 1.47 | 0.75 | 2.02 | 1.07 | 0 | 0.36 | 0.4
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Ag in OREAS 550b (Indicative Value 0.039 ppm) | [18 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 23 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 3
3
| 1 | 1 | < 100 | 0.036 | [15 empty] | 1
| 1 | 2 | < 100 | 0.036 | [15 empty] | 4
| 1 | 3 | < 100 | 0.045 | [15 empty] | 16
| 1 | 4 | < 100 | 0.04 | [15 empty] | 0.0385
| 1 | 5 | < 100 | 0.038 | [15 empty] | 10
| 1 | 6 | < 100 | 0.036
| Mean |  |  | 0.0385
| Median |  |  | 0.037
| Std Dev. |  |  | 0.0035637
| Rel.Std.Dev. |  |  | 9.256%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Al2O3 in OREAS 550b (Certified Value 11.4 wt.%) | [18 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 11.44 | 11.3 | 11.62 | 11.277 | 11.52 | 10.921 | 11.35 | 11.1 | 11.55 | 11.5 | 11.2398 | 11.46 | 11.61 | 13.24 | 9.58 | 11.3 |  | 1
| 1 | 2 | 11.44 | 11.5 | 11.62 | 11.265 | 11.62 | 11.072 | 11.35 | 11.3 | 11.55 | 11.55 | 11.2804 | 11.6 | 11.65 | 13.26 | 9.62 | 11.2 |  | #N/A
| 1 | 3 | 11.43 | 11.47 | 11.44 | 11.26 | 11.61 | 11.186 | 11.4 | 11.4 | 11.6 | 11.55 | 11.291 | 11.58 | 11.61 | 13.32 | 10.3 | 11.3 |  | 16
| 1 | 4 | 11.39 | 11.48 | 11.48 | 11.28 | 11.64 | 11.072 | 11.3 | 11.15 | 11.6 | 11.35 | 11.1548 | 11.56 | 11.64 | 13.26 | 9.67 | 11.3 |  | 11.4037
| 1 | 5 | 11.31 | 11.73 | 11.54 | 11.25 | 11.3 | 11.243 | 11.4 | 11.3 | 11.55 | 11.45 | 11.307 | 11.42 | 11.57 | 13.34 | 10.19 | 11.3 |  | 73
| 1 | 6 | 11.32 | 11.44 | 11.52 | 11.27 | 11.53 | 10.978 | 11.45 | 11.3 | 11.35 | 11.6 | 11.2803 | 11.59 | 11.6 | 13.27 | 10.53 | 11.3
| Mean |  | 11.3883 | 11.4867 | 11.5367 | 11.267 | 11.5367 | 11.0787 | 11.375 | 11.2583 | 11.5333 | 11.5 | 11.2589 | 11.535 | 11.6133 | 13.2817 | 9.9817 | 11.2833
| Median |  | 11.41 | 11.475 | 11.53 | 11.2675 | 11.57 | 11.072 | 11.375 | 11.3 | 11.55 | 11.525 | 11.2803 | 11.57 | 11.61 | 13.265 | 9.93 | 11.3
| Std Dev. |  | 0.059805 | 0.13909 | 0.073121 | 0.011136 | 0.12596 | 0.12133 | 0.05244 | 0.11143 | 0.093095 | 0.089443 | 0.055621 | 0.075829 | 0.028752 | 0.0392 | 0.40858 | 0.040825
| Rel.Std.Dev. |  | 0.5251% | 1.211% | 0.6338% | 0.09883% | 1.092% | 1.095% | 0.461% | 0.9898% | 0.8072% | 0.7778% | 0.494% | 0.6574% | 0.2476% | 0.2951% | 4.093% | 0.3618%
| PDM3 |  | -0.1344% | 0.7278% | 1.166% | -1.198% | 1.166% | -2.85% | -0.2514% | -1.274% | 1.137% | 0.8448% | -1.27% | 1.152% | 1.839% | 16.47% | -12.47% | -1.055%
| Z-Score (Absolute) |  | 0.26 | 0.26 | 0.53 | 0.91 | 0.53 | 1.91 | 0.33 | 0.95 | 0.51 | 0.33 | 0.95 | 0.52 | 0.94 | 9.82 | 7.75 | 0.82
| Analytical results for As in OREAS 550b (Indicative Value 86 ppm) | [18 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 23 | 25 | 26 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | 100 | 100 | 100 | < 100 | < 100 | 54 | < 80 | [10 empty] | 1
| 1 | 2 | 100 | < 100 | 100 | < 100 | < 100 | 51 | < 80 | [10 empty] | 5
| 1 | 3 | 100 | < 100 | 200 | 100 | < 100 | 51 | < 80 | [10 empty] | 16
| 1 | 4 | 100 | < 100 | 200 | 100 | < 100 | 55 | < 80 | [10 empty] | 85.6667
| 1 | 5 | 100 | < 100 | 100 | < 100 | < 100 | 53 | < 80 | [10 empty] | 11
| 1 | 6 | 100 | 100 | 100 | 100 | < 100 | 56 | < 80
| Mean |  | 100 | 100 | 133.3333 | 100 |  | 53.3333
| Median |  | 100 | 100 | 100 | 100 |  | 53.5
| Std Dev. |  | 0 | 0 | 51.6398 | 0 |  | 2.0656
| Rel.Std.Dev. |  | 0% | 0% | 38.73% | 0% |  | 3.873%
| PDM3 |  | 16.73% | 16.73% | 55.64% | 16.73% |  | -37.74%
| Z-Score (Absolute) |  | 1.35 | 0 | 2.7 | 0.34 | 0.67 | 0.54 | 1.17
| Analytical results for BaO in OREAS 550b (Certified Value 824 ppm) | [18 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 22 | 25 | 26 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  |  |  |  | 0
0
| 1 | 1 | 800 | 850 | 700 | < 11 | 800 | 700 | 900 | 900 | 843.8195 | 900 | 700 | 780.4 | 1000 |  |  |  |  | 1
| 1 | 2 | 800 | 800 | 500 | < 11 | 800 | 800 | 800 | 900 | 825.7339 | 900 | 700 | 786 | 900 |  |  |  |  | #N/A
| 1 | 3 | 800 | 800 | 600 | < 11 | 800 | 800 | 900 | 900 | 843.9722 | 800 | 700 | 788.2 | 1000 |  |  |  |  | 16
| 1 | 4 | 800 | 700 | 600 | < 11 | 800 | 800 | 900 | 1000 | 825.9946 | 900 | 700 | 791.6 | 800 |  |  |  |  | 824.1849
| 1 | 5 | 800 | 900 | 600 | < 11 | 800 | 900 | 900 | 900 | 825.764 | 900 | 700 | 793.8 | 800 |  |  |  |  | 74
| 1 | 6 | 800 | 800 | 400 | < 11 | 800 | 800 | 800 | 900 | 843.6085 | 800 | 700 | 797.2 | 800
| Mean |  | 800 | 808.3333 | 566.6667 |  | 800 | 800 | 866.6667 | 916.6667 | 834.8155 | 866.6667 | 700 | 789.5333 | 883.3333
| Median |  | 800 | 800 | 600 |  | 800 | 800 | 900 | 900 | 834.8016 | 900 | 700 | 789.9 | 850
| Std Dev. |  | 0 | 66.458 | 103.2796 |  | 0 | 63.2456 | 51.6398 | 40.8248 | 9.8432 | 51.6398 | 0 | 5.9822 | 98.3192
| Rel.Std.Dev. |  | 0% | 8.222% | 18.23% |  | 0% | 7.906% | 5.958% | 4.454% | 1.179% | 5.958% | 0% | 0.7577% | 11.13%
| PDM3 |  | -2.934% | -1.923% | -31.25% |  | -2.934% | -2.934% | 5.154% | 11.22% | 1.29% | 5.154% | -15.07% | -4.204% | 7.177%
| Z-Score (Absolute) |  | 0 | 0.08 | 2.36 | 8.04 | 0 | 0 | 0.67 | 1.18 | 0.35 | 0.67 | 1.01 | 0.11 | 0.84
| Analytical results for Be in OREAS 550b (Indicative Value 2 ppm) | [18 empty] | Indicative
Be | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 3 | [16 empty] | 1
| 1 | 2 | 2 | [16 empty] | 6
| 1 | 3 | 2 | [16 empty] | 16
| 1 | 4 | 2 | [16 empty] | 2
| 1 | 5 | 1 | [16 empty] | 12
| 1 | 6 | 2
| Mean |  | 2
| Median |  | 2
| Std Dev. |  | 0.63246
| Rel.Std.Dev. |  | 31.62%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 550b (Indicative Value < 100 ppm) | [18 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 23 | 25 | 26 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 10 | < 100 | 200 | < 100 | < 100 | 2.2 | < 90 | [10 empty] | 1
| 1 | 2 | < 10 | < 100 | 200 | < 100 | < 100 | 2 | < 90 | [10 empty] | 7
| 1 | 3 | < 10 | < 100 | 200 | < 100 | < 100 | 2.2 | < 90 | [10 empty] | 16
| 1 | 4 | < 10 | < 100 | 200 | < 100 | < 100 | 2.4 | < 90 | [10 empty] | < 100
| 1 | 5 | < 10 | < 100 | 200 | < 100 | < 100 | 2.1 | < 90 | [10 empty] | 13
| 1 | 6 | < 10 | < 100 | 200 | < 100 | < 100 | 2.1 | < 90
| Mean |  |  |  | 200 |  |  | 2.1667
| Median |  |  |  | 200 |  |  | 2.15
| Std Dev. |  |  |  | 0 |  |  | 0.13663
| Rel.Std.Dev. |  |  |  | 0% |  |  | 6.306%
PDM3
| Z-Score (Absolute) |  | 5.89 | 0 | 19.64 | 0 | 0 | 6.26 | 0.67
| Analytical results for CaO in OREAS 550b (Certified Value 1.87 wt.%) | [18 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 1.88 | 1.86 | 1.88 | 1.867 | 2.06 | 1.819 | 1.88 | 1.8 | 1.88 | 1.9 | 1.8609 | 1.8 | 1.93 | 2.49 | 1.65 | 1.91 |  | 1
| 1 | 2 | 1.89 | 1.88 | 1.89 | 1.87 | 2.07 | 1.833 | 1.88 | 1.88 | 1.86 | 1.9 | 1.8674 | 1.81 | 1.93 | 2.59 | 1.67 | 1.87 |  | #N/A
| 1 | 3 | 1.89 | 1.89 | 1.86 | 1.86 | 2.05 | 1.8469 | 1.87 | 1.86 | 1.88 | 1.9 | 1.8608 | 1.82 | 1.93 | 2.53 | 1.78 | 1.86 |  | 16
| 1 | 4 | 1.89 | 1.87 | 1.87 | 1.88 | 2.02 | 1.833 | 1.86 | 1.86 | 1.88 | 1.9 | 1.8408 | 1.8 | 1.93 | 2.52 | 1.68 | 1.9 |  | 1.8724
| 1 | 5 | 1.88 | 1.88 | 1.87 | 1.87 | 1.97 | 1.8469 | 1.86 | 1.89 | 1.86 | 1.91 | 1.8802 | 1.77 | 1.94 | 2.49 | 1.72 | 1.86 |  | 75
| 1 | 6 | 1.89 | 1.9 | 1.88 | 1.87 | 2 | 1.819 | 1.88 | 1.86 | 1.85 | 1.92 | 1.8532 | 1.79 | 1.93 | 2.61 | 1.81 | 1.86
| Mean |  | 1.8867 | 1.88 | 1.875 | 1.8695 | 2.0283 | 1.833 | 1.8717 | 1.8583 | 1.8683 | 1.905 | 1.8606 | 1.7983 | 1.9317 | 2.5383 | 1.7183 | 1.8767
| Median |  | 1.89 | 1.88 | 1.875 | 1.87 | 2.035 | 1.833 | 1.875 | 1.86 | 1.87 | 1.9 | 1.8609 | 1.8 | 1.93 | 2.525 | 1.7 | 1.865
| Std Dev. |  | 0.005164 | 0.014142 | 0.010488 | 0.006442 | 0.038687 | 0.012477 | 0.0098319 | 0.031252 | 0.013292 | 0.0083666 | 0.013251 | 0.017224 | 0.0040825 | 0.050761 | 0.064317 | 0.022509
| Rel.Std.Dev. |  | 0.2737% | 0.7522% | 0.5594% | 0.3446% | 1.907% | 0.6807% | 0.5253% | 1.682% | 0.7114% | 0.4392% | 0.7122% | 0.9578% | 0.2113% | 2% | 3.743% | 1.199%
| PDM3 |  | 0.7622% | 0.4061% | 0.1391% | -0.1546% | 8.328% | -2.106% | -0.03893% | -0.751% | -0.217% | 1.741% | -0.6318% | -3.955% | 3.166% | 35.57% | -8.228% | 0.2281%
| Z-Score (Absolute) |  | 0.63 | 0.32 | 0.08 | 0.18 | 7.38 | 1.92 | 0.08 | 0.71 | 0.24 | 1.51 | 0.61 | 3.57 | 2.78 | 31.65 | 7.38 | 0.16
| Analytical results for Cd in OREAS 550b (Indicative Value < 10 ppm) | [18 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 07 | 23 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | < 100 | [15 empty] | 1
| 1 | 2 | < 10 | < 100 | [15 empty] | 8
| 1 | 3 | < 10 | < 100 | [15 empty] | 16
| 1 | 4 | < 10 | < 100 | [15 empty] | < 10
| 1 | 5 | < 10 | < 100 | [15 empty] | 14
| 1 | 6 | < 10 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 550b (Indicative Value 123 ppm) | [18 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 160 | 160 | 125.16 | 122 | 260 | [12 empty] | 1
| 1 | 2 | 80 | 80 | 124.96 | 115 | < 170 | [12 empty] | 9
| 1 | 3 | 160 | 160 | 125.08 | 123 | < 170 | [12 empty] | 16
| 1 | 4 | 80 | 80 | 125.19 | 122 | 170 | [12 empty] | 122.5779
| 1 | 5 | 160 | 160 | 125.17 | 118 | 170 | [12 empty] | 15
| 1 | 6 | 80 | 80 | 124.99 | 126 | < 170
| Mean |  | 120 | 120 | 125.0917 | 121 | 200
| Median |  | 120 | 120 | 125.12 | 122 | 170
| Std Dev. |  | 43.8178 | 43.8178 | 0.098268 | 3.8987 | 51.9615
| Rel.Std.Dev. |  | 36.51% | 36.51% | 0.07856% | 3.222% | 25.98%
| PDM3 |  | -2.103% | -2.103% | 2.051% | -1.287% | 63.16%
| Z-Score (Absolute) |  | 0 | 0 | 1.81 | 0.67 | 12.26
| Analytical results for Cl in OREAS 550b (Indicative Value 350 ppm) | [18 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 04 | 07 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 20 | 350 | [15 empty] | 1
| 1 | 2 | < 20 | 350 | [15 empty] | 10
| 1 | 3 | < 20 | 350 | [15 empty] | 16
| 1 | 4 | < 20 | 350 | [15 empty] | 350
| 1 | 5 | < 20 | 350 | [15 empty] | 16
| 1 | 6 | < 20 | 350
| Mean |  |  | 350
| Median |  |  | 350
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 550b (Certified Value 0.145 wt.%) | [18 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 13 | 14 | 19 | 20 | 22 | 23 | 25 | 26 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  |  | 3
3
| 1 | 1 | 0.144 | 0.1468 | 0.149 | 0.147 | 0.14 | 0.14 | 0.14 | 0.15 | 0.14 | 0.15 | 0.145 | 0.12 | 0.1422 | 0.14019 | 0.154 |  |  | 1
| 1 | 2 | 0.144 | 0.1476 | 0.15 | 0.147 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.15 | 0.144 | 0.14 | 0.1397 | 0.13903 | 0.154 |  |  | 29
| 1 | 3 | 0.143 | 0.1498 | 0.146 | 0.147 | 0.14 | 0.14 | 0.15 | 0.15 | 0.14 | 0.15 | 0.145 | 0.11 | 0.14 | 0.1397 | 0.147 |  |  | 16
| 1 | 4 | 0.146 | 0.1476 | 0.149 | 0.147 | 0.14 | 0.14 | 0.14 | 0.15 | 0.14 | 0.15 | 0.145 | 0.13 | 0.1378 | 0.13851 | 0.154 |  |  | 0.14455
| 1 | 5 | 0.144 | 0.149 | 0.146 | 0.147 | 0.14 | 0.14 | 0.15 | 0.15 | 0.14 | 0.14 | 0.146 | 0.12 | 0.1425 | 0.13817 | 0.147 |  |  | 76
| 1 | 6 | 0.144 | 0.1461 | 0.149 | 0.147 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14 | 0.15 | 0.145 | 0.14 | 0.1432 | 0.13883 | 0.154
| Mean |  | 0.14417 | 0.14782 | 0.14817 | 0.147 | 0.14 | 0.14 | 0.145 | 0.14667 | 0.14 | 0.14833 | 0.145 | 0.12667 | 0.1409 | 0.13907 | 0.15167
| Median |  | 0.144 | 0.1476 | 0.149 | 0.147 | 0.14 | 0.14 | 0.145 | 0.15 | 0.14 | 0.15 | 0.145 | 0.125 | 0.1411 | 0.13893 | 0.154
| Std Dev. |  | 0.00098319 | 0.0013717 | 0.0017224 | 0 | 0 | 0 | 0.0054772 | 0.005164 | 0 | 0.0040825 | 0.00063246 | 0.012111 | 0.0020688 | 0.00075244 | 0.0036148
| Rel.Std.Dev. |  | 0.682% | 0.928% | 1.162% | 0% | 0% | 0% | 3.777% | 3.521% | 0% | 2.752% | 0.4362% | 9.561% | 1.468% | 0.541% | 2.383%
| PDM3 |  | -0.2659% | 2.259% | 2.501% | 1.694% | -3.148% | -3.148% | 0.3106% | 1.464% | -3.148% | 2.617% | 0.3106% | -12.37% | -2.526% | -3.791% | 4.923%
| Z-Score (Absolute) |  | 0.17 | 0.57 | 0.64 | 0.37 | 1.01 | 1.01 | 0 | 0.34 | 1.01 | 0.67 | 0 | 3.71 | 0.83 | 1.2 | 1.36
| Analytical results for Cr2O3 in OREAS 550b (Indicative Value 118 ppm) | [18 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  |  | 0
0
| 1 | 1 | 100 | 100 | 60 | < 100 | < 15 | 100 | 100 | 150 | 150 | 136.0999 | 60 | 100 | < 100 | 90.6 | 200 |  |  | 1
| 1 | 2 | 100 | 100 | 90 | < 100 | < 15 | < 100 | 200 | 440 | 150 | 145.184 | < 50 | 100 | < 100 | 102.3 | 200 |  |  | 11
| 1 | 3 | 100 | 100 | 50 | < 100 | < 15 | 100 | 100 | 290 | 150 | 136.1243 | 90 | 90 | < 100 | 90.6 | < 100 |  |  | 16
| 1 | 4 | 100 | 100 | 80 | 200 | < 15 | < 100 | 100 | 150 | 150 | 154.3067 | 50 | 90 | < 100 | 92.1 | 100 |  |  | 117.533
| 1 | 5 | 100 | 200 | 60 | < 100 | < 15 | < 100 | 100 | 290 | 150 | 145.1893 | < 50 | 100 | < 100 | 87.7 | < 100 |  |  | 17
| 1 | 6 | 100 | 100 | < 50 | < 100 | < 15 | 100 | 100 | 440 | 150 | 145.137 | < 50 | 90 | < 100 | 93.5 | < 100
| Mean |  | 100 | 116.6667 | 68 | 200 |  | 100 | 116.6667 | 293.3333 | 150 | 143.6735 | 66.6667 | 95 |  | 92.8 | 166.6667
| Median |  | 100 | 100 | 60 | 200 |  | 100 | 100 | 290 | 150 | 145.1605 | 60 | 95 |  | 91.35 | 200
| Std Dev. |  | 0 | 40.8248 | 16.4317 |  |  | 0 | 40.8248 | 129.7176 | 0 | 6.843 | 20.8167 | 5.4772 |  | 5.0383 | 57.735
| Rel.Std.Dev. |  | 0% | 34.99% | 24.16% |  |  | 0% | 34.99% | 44.22% | 0% | 4.763% | 31.22% | 5.766% |  | 5.429% | 34.64%
| PDM3 |  | -14.92% | -0.7371% | -42.14% | 70.17% |  | -14.92% | -0.7371% | 149.6% | 27.62% | 22.24% | -43.28% | -19.17% |  | -21.04% | 41.8%
| Z-Score (Absolute) |  | 0.16 | 0.67 | 1.06 | 0.62 | 2.73 | 0.62 | 0.67 | 6.14 | 1.59 | 1.51 | 1.53 | 0 | 1.4 | 0.07 | 0.41
| Analytical results for Cs in OREAS 550b (Indicative Value 2.12 ppm) | [18 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 25 | 26 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [15 empty] | 2
2
| 1 | 1 | 2.1 | < 90 | [15 empty] | 1
| 1 | 2 | 2.1 | < 90 | [15 empty] | 12
| 1 | 3 | 2.1 | < 90 | [15 empty] | 16
| 1 | 4 | 2 | < 90 | [15 empty] | 2.1167
| 1 | 5 | 2.2 | < 90 | [15 empty] | 18
| 1 | 6 | 2.2 | < 90
| Mean |  | 2.1167
| Median |  | 2.1
| Std Dev. |  | 0.075277
| Rel.Std.Dev. |  | 3.556%
| PDM3 |  | -1.443e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 550b (Certified Value 1.84 wt.%) | [18 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 1.831 | 1.8054 | 1.813 | 1.8221 | 2.149 | 1.82 | 1.815 | 1.825 | 1.815 | 1.815 | 1.8873 | 1.85 | 1.816 | 1.64 | 2.0772 | 2.009 | 1.909 | 1
| 1 | 2 | 1.82 | 1.8214 | 1.867 | 1.8266 | 2.069 | 1.83 | 1.845 | 1.865 | 1.815 | 1.81 | 1.8946 | 1.87 | 1.812 | 1.65 | 2.0053 | 2.0535 | 1.877 | 30
| 1 | 3 | 1.81 | 1.8453 | 1.849 | 1.8134 | 2.165 | 1.86 | 1.85 | 1.87 | 1.8 | 1.855 | 1.8958 | 1.86 | 1.833 | 1.65 | 1.9573 | 2.0439 | 1.853 | 16
| 1 | 4 | 1.822 | 1.8134 | 1.911 | 1.8294 | 2.173 | 1.81 | 1.825 | 1.845 | 1.81 | 1.81 | 1.8724 | 1.88 | 1.834 | 1.63 | 2.0372 | 2.0835 | 1.885 | 1.8423
| 1 | 5 | 1.833 | 1.8453 | 1.893 | 1.8214 | 1.989 | 1.83 | 1.835 | 1.88 | 1.795 | 1.81 | 1.8918 | 1.83 | 1.824 | 1.64 | 2.0772 | 2.0294 | 1.917 | 77
| 1 | 6 | 1.837 | 1.8374 | 1.847 | 1.8214 | 2.013 | 1.84 | 1.855 | 1.865 | 1.81 | 1.82 | 1.884 | 1.87 | 1.8 | 1.64 | 2.0852 | 2.0233 | 1.885
| Mean |  | 1.8255 | 1.828 | 1.8633 | 1.8224 | 2.093 | 1.8317 | 1.8375 | 1.8583 | 1.8075 | 1.82 | 1.8877 | 1.86 | 1.8198 | 1.6417 | 2.0399 | 2.0404 | 1.8877
| Median |  | 1.8265 | 1.8294 | 1.858 | 1.8218 | 2.109 | 1.83 | 1.84 | 1.865 | 1.81 | 1.8125 | 1.8896 | 1.865 | 1.82 | 1.64 | 2.0572 | 2.0366 | 1.885
| Std Dev. |  | 0.010015 | 0.017055 | 0.035132 | 0.0054683 | 0.080637 | 0.017224 | 0.015411 | 0.019916 | 0.0082158 | 0.017607 | 0.0086847 | 0.017889 | 0.013121 | 0.0075277 | 0.050696 | 0.026242 | 0.023001
| Rel.Std.Dev. |  | 0.5486% | 0.933% | 1.885% | 0.3001% | 3.853% | 0.9403% | 0.8387% | 1.072% | 0.4545% | 0.9674% | 0.4601% | 0.9617% | 0.721% | 0.4585% | 2.485% | 1.286% | 1.219%
| PDM3 |  | -0.9111% | -0.7735% | 1.143% | -1.081% | 13.61% | -0.5763% | -0.2597% | 0.8712% | -1.888% | -1.21% | 2.464% | 0.9616% | -1.219% | -10.89% | 10.73% | 10.76% | 2.463%
| Z-Score (Absolute) |  | 0.36 | 0.28 | 0.77 | 0.45 | 7.66 | 0.17 | 0 | 0.62 | 0.9 | 0.52 | 1.5 | 0.67 | 0.53 | 5.87 | 6.07 | 6.08 | 1.51
| Analytical results for Dy in OREAS 550b (Indicative Value 5.28 ppm) | [18 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 5.44 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 5.04 | < 170 | [14 empty] | 13
| 1 | 3 | < 10 | 5.38 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 5.18 | < 170 | [14 empty] | 5.2817
| 1 | 5 | < 10 | 4.99 | < 170 | [14 empty] | 19
| 1 | 6 | < 10 | 5.66 | < 170
| Mean |  |  | 5.2817
| Median |  |  | 5.28
| Std Dev. |  |  | 0.25756
| Rel.Std.Dev. |  |  | 4.876%
| PDM3 |  |  | -5.551e-14%
| Z-Score (Absolute) |  | 0.67 | 0 | 195.95
| Analytical results for Er in OREAS 550b (Indicative Value 2.99 ppm) | [18 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 3.12 | [15 empty] | 1
| 1 | 2 | < 10 | 2.83 | [15 empty] | 14
| 1 | 3 | < 10 | 2.94 | [15 empty] | 16
| 1 | 4 | < 10 | 2.98 | [15 empty] | 2.99
| 1 | 5 | < 10 | 2.93 | [15 empty] | 20
| 1 | 6 | < 10 | 3.14
| Mean |  |  | 2.99
| Median |  |  | 2.96
| Std Dev. |  |  | 0.11933
| Rel.Std.Dev. |  |  | 3.991%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 550b (Indicative Value 0.86 ppm) | [18 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 0.85 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 0.77 | < 170 | [14 empty] | 15
| 1 | 3 | < 10 | 0.86 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 0.87 | < 170 | [14 empty] | 0.85667
| 1 | 5 | < 10 | 0.88 | < 170 | [14 empty] | 21
| 1 | 6 | < 10 | 0.91 | < 170
| Mean |  |  | 0.85667
| Median |  |  | 0.865
| Std Dev. |  |  | 0.047188
| Rel.Std.Dev. |  |  | 5.508%
| PDM3 |  |  | -2.22e-14%
| Z-Score (Absolute) |  | 0 | 0.67 | 13.24
| Analytical results for F in OREAS 550b (Indicative Value < 5000 ppm) | [18 empty] | Indicative
F | Round | Replicate | Lab | [16 empty] | 1
| No | No | 26 | [16 empty] | ppm
|  |  | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 5000 | [16 empty] | 1
| 1 | 2 | < 5000 | [16 empty] | 16
| 1 | 3 | < 5000 | [16 empty] | 16
| 1 | 4 | < 5000 | [16 empty] | < 5000
| 1 | 5 | < 5000 | [16 empty] | 22
| 1 | 6 | < 5000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 550b (Certified Value 4.01 wt.%) | [18 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 3.99 | 4 | 4.15 | 4.001 | 4.003 | 3.975 | 3.917 | 3.84 | 3.989 | 3.932 | 3.9776 | 4 | 3.92 | 4.09 | 4.0197 | 4.13 | 4.01 | 1
| 1 | 2 | 4 | 4.05 | 4.18 | 4.014 | 4.046 | 3.989 | 3.917 | 3.92 | 3.975 | 3.932 | 3.9907 | 4.07 | 3.94 | 4.13 | 3.9597 | 4 | 3.93 | #N/A
| 1 | 3 | 3.99 | 4.07 | 4.12 | 4.02 | 4.017 | 4.046 | 3.917 | 3.96 | 4.032 | 3.975 | 3.9852 | 4.04 | 3.94 | 4.12 | 4.0797 | 4.44 | 3.98 | 16
| 1 | 4 | 4 | 4.03 | 4.14 | 4.01 | 4.003 | 4.032 | 3.889 | 3.9 | 3.989 | 3.96 | 3.9376 | 4.06 | 3.96 | 4.07 | 4.1097 | 3.66 | 4 | 4.0116
| 1 | 5 | 3.96 | 4.08 | 4.13 | 4.02 | 3.932 | 4.003 | 3.932 | 3.98 | 3.989 | 3.975 | 4.0056 | 4 | 3.98 | 4.15 | 4.1097 | 4.19 | 4.24 | 78
| 1 | 6 | 3.95 | 4.03 | 4.14 | 4 | 3.989 | 3.946 | 3.975 | 3.97 | 3.946 | 3.96 | 3.9756 | 4.03 | 3.97 | 4.12 | 4.0797 | 4.14 | 3.95
| Mean |  | 3.9817 | 4.0433 | 4.1433 | 4.0108 | 3.9983 | 3.9985 | 3.9245 | 3.9283 | 3.9867 | 3.9557 | 3.9787 | 4.0333 | 3.9517 | 4.1133 | 4.0597 | 4.0933 | 4.0183
| Median |  | 3.99 | 4.04 | 4.14 | 4.012 | 4.003 | 3.996 | 3.917 | 3.94 | 3.989 | 3.96 | 3.9814 | 4.035 | 3.95 | 4.12 | 4.0797 | 4.135 | 3.99
| Std Dev. |  | 0.02137 | 0.029439 | 0.020656 | 0.0088638 | 0.037819 | 0.036871 | 0.02841 | 0.053072 | 0.02779 | 0.019521 | 0.022839 | 0.029439 | 0.022286 | 0.028752 | 0.058987 | 0.25672 | 0.11268
| Rel.Std.Dev. |  | 0.5367% | 0.7281% | 0.4985% | 0.221% | 0.9459% | 0.9221% | 0.7239% | 1.351% | 0.6971% | 0.4935% | 0.574% | 0.7299% | 0.564% | 0.699% | 1.453% | 6.272% | 2.804%
| PDM3 |  | -0.7453% | 0.792% | 3.285% | -0.0182% | -0.3298% | -0.3256% | -2.17% | -2.075% | -0.6206% | -1.393% | -0.8185% | 0.5427% | -1.493% | 2.537% | 1.2% | 2.038% | 0.1688%
| Z-Score (Absolute) |  | 0.26 | 0.71 | 2.28 | 0.2 | 0 | 0 | 1.16 | 1.1 | 0.19 | 0.67 | 0.31 | 0.55 | 0.73 | 1.81 | 0.96 | 1.49 | 0.31
| Analytical results for Ga in OREAS 550b (Indicative Value 22.3 ppm) | [18 empty] | Indicative
Ga | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 1
1
| 1 | 1 | 21.6 | [16 empty] | 1
| 1 | 2 | 21.9 | [16 empty] | 17
| 1 | 3 | 21.9 | [16 empty] | 16
| 1 | 4 | 21.8 | [16 empty] | 22.3
| 1 | 5 | 24 | [16 empty] | 23
| 1 | 6 | 22.6
| Mean |  | 22.3
| Median |  | 21.9
| Std Dev. |  | 0.89889
| Rel.Std.Dev. |  | 4.031%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 550b (Indicative Value 5.83 ppm) | [18 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 5.8 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 5.82 | < 170 | [14 empty] | 18
| 1 | 3 | < 10 | 5.91 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 5.77 | < 170 | [14 empty] | 5.825
| 1 | 5 | < 10 | 5.64 | < 170 | [14 empty] | 24
| 1 | 6 | < 10 | 6.01 | < 170
| Mean |  |  | 5.825
| Median |  |  | 5.81
| Std Dev. |  |  | 0.12598
| Rel.Std.Dev. |  |  | 2.163%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0 | 66.15
| Analytical results for Hf in OREAS 550b (Indicative Value < 80 ppm) | [18 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 80 | < 80 | < 10 | 5 | 80 | [12 empty] | 1
| 1 | 2 | < 80 | < 80 | < 10 | 4.7 | < 80 | [12 empty] | 22
| 1 | 3 | < 80 | < 80 | < 10 | 4.7 | < 80 | [12 empty] | 16
| 1 | 4 | < 80 | < 80 | < 10 | 4.9 | 80 | [12 empty] | < 80
| 1 | 5 | < 80 | < 80 | < 10 | 4.6 | < 80 | [12 empty] | 10
| 1 | 6 | < 80 | < 80 | < 10 | 4.9 | < 80
| Mean |  |  |  |  | 4.8 | 80
| Median |  |  |  |  | 4.8 | 80
| Std Dev. |  |  |  |  | 0.15492 | 0
| Rel.Std.Dev. |  |  |  |  | 3.227% | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 1.78 | 1.79 | 0.67
| Analytical results for Ho in OREAS 550b (Indicative Value 1.08 ppm) | [18 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 1.08 | [15 empty] | 1
| 1 | 2 | < 10 | 1.06 | [15 empty] | 5
| 1 | 3 | < 10 | 1.08 | [15 empty] | 16
| 1 | 4 | < 10 | 1.1 | [15 empty] | 1.08
| 1 | 5 | < 10 | 1.01 | [15 empty] | 11
| 1 | 6 | < 10 | 1.15
| Mean |  |  | 1.08
| Median |  |  | 1.08
| Std Dev. |  |  | 0.046043
| Rel.Std.Dev. |  |  | 4.263%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 550b (Indicative Value 0.98 ppm) | [18 empty] | Indicative
In | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 0.9 | [16 empty] | 1
| 1 | 2 | 0.9 | [16 empty] | 6
| 1 | 3 | 0.9 | [16 empty] | 16
| 1 | 4 | 1.1 | [16 empty] | 0.98333
| 1 | 5 | 1 | [16 empty] | 12
| 1 | 6 | 1.1
| Mean |  | 0.98333
| Median |  | 0.95
| Std Dev. |  | 0.098319
| Rel.Std.Dev. |  | 9.999%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 550b (Certified Value 2.42 wt.%) | [18 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 2.4 | 2.39 | 2.4 | 2.439 | 2.51 | 2.3249 | 2.45 | 2.32 | 2.44 | 2.46 | 2.4217 | 2.427 | 2.438 | 3.2 | 2.27 | 2.39 |  | 1
| 1 | 2 | 2.39 | 2.42 | 2.41 | 2.441 | 2.53 | 2.349 | 2.47 | 2.39 | 2.44 | 2.47 | 2.4309 | 2.451 | 2.446 | 3.2 | 2.13 | 2.36 |  | #N/A
| 1 | 3 | 2.42 | 2.46 | 2.37 | 2.43 | 2.48 | 2.349 | 2.46 | 2.43 | 2.45 | 2.45 | 2.4316 | 2.434 | 2.438 | 3.13 | 2.27 | 2.39 |  | 16
| 1 | 4 | 2.41 | 2.42 | 2.38 | 2.45 | 2.45 | 2.3369 | 2.47 | 2.39 | 2.45 | 2.47 | 2.4008 | 2.434 | 2.443 | 3.29 | 2.18 | 2.39 |  | 2.4195
| 1 | 5 | 2.41 | 2.41 | 2.38 | 2.44 | 2.37 | 2.3249 | 2.42 | 2.43 | 2.44 | 2.47 | 2.4392 | 2.388 | 2.441 | 3.27 | 2.15 | 2.39 |  | 79
| 1 | 6 | 2.41 | 2.4 | 2.39 | 2.45 | 2.42 | 2.3128 | 2.45 | 2.42 | 2.41 | 2.47 | 2.4283 | 2.42 | 2.42 | 3.27 | 2.17 | 2.39
| Mean |  | 2.4067 | 2.4167 | 2.3883 | 2.4417 | 2.46 | 2.3329 | 2.4533 | 2.3967 | 2.4383 | 2.465 | 2.4254 | 2.4257 | 2.4377 | 3.2267 | 2.195 | 2.385
| Median |  | 2.41 | 2.415 | 2.385 | 2.4405 | 2.465 | 2.3309 | 2.455 | 2.405 | 2.44 | 2.47 | 2.4296 | 2.4305 | 2.4395 | 3.235 | 2.175 | 2.39
| Std Dev. |  | 0.010328 | 0.024221 | 0.01472 | 0.0075542 | 0.05933 | 0.014604 | 0.018619 | 0.041793 | 0.01472 | 0.0083666 | 0.013307 | 0.021134 | 0.0091797 | 0.060882 | 0.060581 | 0.012247
| Rel.Std.Dev. |  | 0.4291% | 1.002% | 0.6163% | 0.3094% | 2.412% | 0.626% | 0.7589% | 1.744% | 0.6037% | 0.3394% | 0.5486% | 0.8713% | 0.3766% | 1.887% | 2.76% | 0.5135%
| PDM3 |  | -0.5314% | -0.1181% | -1.289% | 0.9152% | 1.673% | -3.58% | 1.397% | -0.9447% | 0.7774% | 1.88% | 0.244% | 0.2539% | 0.7498% | 33.36% | -9.28% | -1.427%
| Z-Score (Absolute) |  | 0.45 | 0.21 | 0.89 | 0.38 | 0.82 | 2.2 | 0.66 | 0.69 | 0.3 | 0.94 | 0 | 0 | 0.29 | 19.07 | 5.49 | 0.96
| Analytical results for La in OREAS 550b (Indicative Value 73 ppm) | [18 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 90 | 90 | 64.15 | 61.9 | < 170 | [12 empty] | 1
| 1 | 2 | 90 | 90 | 66.48 | 60.1 | < 170 | [12 empty] | 7
| 1 | 3 | 170 | 90 | 64.94 | 62.3 | < 170 | [12 empty] | 16
| 1 | 4 | 90 | 90 | 65.24 | 62.6 | < 170 | [12 empty] | 72.8341
| 1 | 5 | 90 | < 90 | 67.49 | 60.5 | < 170 | [12 empty] | 13
| 1 | 6 | 90 | 90 | 66.16 | 65.6 | < 170
| Mean |  | 103.3333 | 90 | 65.7433 | 62.1667
| Median |  | 90 | 90 | 65.7 | 62.1
| Std Dev. |  | 32.6599 | 0 | 1.2004 | 1.9552
| Rel.Std.Dev. |  | 31.61% | 0% | 1.826% | 3.145%
| PDM3 |  | 41.87% | 23.57% | -9.736% | -14.65%
| Z-Score (Absolute) |  | 1.16 | 0 | 0.67 | 0.87 | 0.39
| Analytical results for Lu in OREAS 550b (Indicative Value 0.42 ppm) | [18 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 0.38 | [15 empty] | 1
| 1 | 2 | < 10 | 0.44 | [15 empty] | 8
| 1 | 3 | < 10 | 0.43 | [15 empty] | 16
| 1 | 4 | < 10 | 0.43 | [15 empty] | 0.41833
| 1 | 5 | < 10 | 0.39 | [15 empty] | 14
| 1 | 6 | < 10 | 0.44
| Mean |  |  | 0.41833
| Median |  |  | 0.43
| Std Dev. |  |  | 0.026394
| Rel.Std.Dev. |  |  | 6.309%
| PDM3 |  |  | 6.661e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 550b (Certified Value 7.19 wt.%) | [18 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 7.18 | 6.99 | 7.33 | 6.986 | 7.04 | 6.981 | 7.22 | 6.89 | 7.41 | 7.26 | 7.2587 | 7.23 | 7.44 | 6.92 | 6.6 | 7.3 |  | 1
| 1 | 2 | 7.22 | 7.04 | 7.36 | 6.978 | 7.25 | 7.064 | 7.25 | 7.06 | 7.37 | 7.3 | 7.2846 | 7.27 | 7.46 | 6.77 | 6.43 | 7.2 |  | #N/A
| 1 | 3 | 7.21 | 7.13 | 7.3 | 6.97 | 7.3 | 7.147 | 7.23 | 7.13 | 7.4 | 7.3 | 7.3004 | 7.23 | 7.44 | 6.96 | 6.64 | 7.2 |  | 16
| 1 | 4 | 7.18 | 7.08 | 7.29 | 6.98 | 7.32 | 7.064 | 7.17 | 6.98 | 7.4 | 7.28 | 7.2134 | 7.26 | 7.44 | 6.99 | 6.49 | 7.2 |  | 7.1901
| 1 | 5 | 7.23 | 7.14 | 7.3 | 6.97 | 7.2 | 7.081 | 7.2 | 7.14 | 7.4 | 7.29 | 7.3248 | 7.1 | 7.42 | 6.79 | 6.44 | 7.2 |  | 80
| 1 | 6 | 7.18 | 7.1 | 7.32 | 6.99 | 7.34 | 7.014 | 7.26 | 7.02 | 7.31 | 7.31 | 7.2886 | 7.18 | 7.44 | 6.96 | 6.76 | 7.2
| Mean |  | 7.2 | 7.08 | 7.3167 | 6.979 | 7.2417 | 7.0585 | 7.2217 | 7.0367 | 7.3817 | 7.29 | 7.2784 | 7.2117 | 7.44 | 6.8983 | 6.56 | 7.2167
| Median |  | 7.195 | 7.09 | 7.31 | 6.979 | 7.275 | 7.064 | 7.225 | 7.04 | 7.4 | 7.295 | 7.2866 | 7.23 | 7.44 | 6.94 | 6.545 | 7.2
| Std Dev. |  | 0.022804 | 0.056921 | 0.02582 | 0.0081731 | 0.11107 | 0.057246 | 0.033116 | 0.094798 | 0.037639 | 0.017889 | 0.038455 | 0.063061 | 0.012649 | 0.094534 | 0.12977 | 0.040825
| Rel.Std.Dev. |  | 0.3167% | 0.804% | 0.3529% | 0.1171% | 1.534% | 0.811% | 0.4586% | 1.347% | 0.5099% | 0.2454% | 0.5283% | 0.8744% | 0.17% | 1.37% | 1.978% | 0.5657%
| PDM3 |  | 0.1383% | -1.531% | 1.761% | -2.935% | 0.7178% | -1.83% | 0.4397% | -2.133% | 2.665% | 1.39% | 1.229% | 0.3006% | 3.476% | -4.057% | -8.763% | 0.3701%
| Z-Score (Absolute) |  | 0.08 | 0.76 | 0.58 | 1.34 | 0.16 | 0.89 | 0.04 | 1.01 | 0.95 | 0.43 | 0.37 | 0.01 | 1.29 | 1.8 | 3.73 | 0.01
| Analytical results for MnO in OREAS 550b (Certified Value 0.026 wt.%) | [18 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.03 | 0.027 | 0.02 | 0.026 | < 0.0013 | 0.013 | 0.02 | 0.026 | 0.013 | 0.021776 | 0.021 | 0.023 | 0.03 | 0.02604 | 0.024 | 0.028 |  | 1
| 1 | 2 | 0.03 | 0.028 | 0.02 | 0.026 | < 0.0013 | 0.026 | 0.03 | 0.026 | 0.013 | 0.022685 | 0.026 | 0.024 | 0.03 | 0.02883 | 0.025 | 0.028 |  | #N/A
| 1 | 3 | 0.03 | 0.027 | 0.02 | 0.026 | < 0.0013 | 0.026 | 0.03 | 0.026 | 0.013 | 0.020873 | 0.02 | 0.023 | 0.03 | 0.02697 | 0.025 | 0.019 |  | 16
| 1 | 4 | 0.03 | 0.027 | 0.02 | 0.026 | < 0.0013 | 0.026 | 0.02 | 0.026 | 0.013 | 0.020877 | 0.02 | 0.023 | 0.03 | 0.0279 | 0.024 | 0.028 |  | 0.025504
| 1 | 5 | 0.03 | 0.029 | 0.02 | 0.026 | < 0.0013 | 0.026 | 0.03 | 0.026 | 0.013 | 0.022686 | 0.021 | 0.024 | 0.04 | 0.02604 | 0.025 | 0.028 |  | 81
| 1 | 6 | 0.03 | 0.028 | 0.02 | 0.026 | < 0.0013 | 0.026 | 0.03 | 0.026 | 0.013 | 0.021771 | 0.022 | 0.024 | 0.03 | 0.0279 | 0.025 | 0.028
| Mean |  | 0.03 | 0.027667 | 0.02 | 0.026 |  | 0.023833 | 0.026667 | 0.026 | 0.013 | 0.021778 | 0.021667 | 0.0235 | 0.031667 | 0.02728 | 0.024667 | 0.0265
| Median |  | 0.03 | 0.0275 | 0.02 | 0.026 |  | 0.026 | 0.03 | 0.026 | 0.013 | 0.021773 | 0.021 | 0.0235 | 0.03 | 0.027435 | 0.025 | 0.028
| Std Dev. |  | 0 | 0.0008165 | 0 | 0 |  | 0.0053072 | 0.005164 | 0 | 0 | 0.0008098 | 0.0022509 | 0.00054772 | 0.0040825 | 0.0011263 | 0.0005164 | 0.0036742
| Rel.Std.Dev. |  | 0% | 2.951% | 0% | 0% |  | 22.27% | 19.36% | 0% | 0% | 3.718% | 10.39% | 2.331% | 12.89% | 4.129% | 2.094% | 13.87%
| PDM3 |  | 17.63% | 8.481% | -21.58% | 1.946% |  | -6.549% | 4.56% | 1.946% | -49.03% | -14.61% | -15.04% | -7.856% | 24.17% | 6.967% | -3.282% | 3.907%
| Z-Score (Absolute) |  | 1.44 | 0.73 | 1.59 | 0.18 | 7.45 | 0.48 | 0.43 | 0.18 | 3.73 | 1.05 | 1.08 | 0.53 | 1.94 | 0.62 | 0.18 | 0.33
| Analytical results for Mo in OREAS 550b (Indicative Value 16.8 ppm) | [18 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 25 | 26 | [11 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | 27 | < 50 | < 50 | < 50 | 6.4 | 400 | [11 empty] | 1
| 1 | 2 | 27 | < 50 | < 50 | < 50 | 6.6 | 330 | [11 empty] | 9
| 1 | 3 | 27 | < 50 | < 50 | < 50 | 6.8 | 270 | [11 empty] | 16
| 1 | 4 | 27 | < 50 | < 50 | < 50 | 6.6 | 470 | [11 empty] | 16.7667
| 1 | 5 | 26 | < 50 | < 50 | < 50 | 6.9 | 400 | [11 empty] | 15
| 1 | 6 | 27 | < 50 | < 50 | < 50 | 6.9 | 400
| Mean |  | 26.8333 |  |  |  | 6.7 | 378.3333
| Median |  | 27 |  |  |  | 6.7 | 400
| Std Dev. |  | 0.40825 |  |  |  | 0.2 | 69.1134
| Rel.Std.Dev. |  | 1.521% |  |  |  | 2.985% | 18.27%
| PDM3 |  | 60.04% |  |  |  | -60.04% | 2156%
| Z-Score (Absolute) |  | 1.35 | 0 | 0 | 0 | 13.46 | 259.45
| Analytical results for Na2O in OREAS 550b (Indicative Value 0.084 wt.%) | [18 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 15 | 19 | 22 | 25 | 26 | [8 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 3
3
| 1 | 1 | 0.08 | 0.06 | 0.08 | < 0.01 | < 0.01 | 0.1 | 0.13 | 0.05 | 0.1 | [8 empty] | 1
| 1 | 2 | 0.09 | 0.07 | 0.07 | < 0.01 | < 0.01 | 0.08 | 0.13 | 0.05 | 0.1 | [8 empty] | 10
| 1 | 3 | 0.08 | 0.06 | 0.08 | < 0.01 | < 0.01 | 0.06 | 0.13 | 0.06 | < 0.1 | [8 empty] | 16
| 1 | 4 | 0.1 | 0.05 | 0.09 | < 0.01 | < 0.01 | 0.05 | 0.13 | 0.05 | 0.1 | [8 empty] | 0.083571
| 1 | 5 | 0.1 | 0.05 | 0.08 | < 0.01 | < 0.01 | 0.05 | 0.13 | 0.07 | < 0.1 | [8 empty] | 16
| 1 | 6 | 0.09 | 0.06 | 0.09 | < 0.01 | < 0.01 | 0.06 | 0.13 | 0.07 | < 0.1
| Mean |  | 0.09 | 0.058333 | 0.081667 |  |  | 0.066667 | 0.13 | 0.058333 | 0.1
| Median |  | 0.09 | 0.06 | 0.08 |  |  | 0.06 | 0.13 | 0.055 | 0.1
| Std Dev. |  | 0.0089443 | 0.0075277 | 0.0075277 |  |  | 0.019664 | 0 | 0.0098319 | 1.6997e-17
| Rel.Std.Dev. |  | 9.938% | 12.9% | 9.218% |  |  | 29.5% | 0% | 16.85% | 1.7e-14%
| PDM3 |  | 7.692% | -30.2% | -2.279% |  |  | -20.23% | 55.56% | -30.2% | 19.66%
| Z-Score (Absolute) |  | 1.05 | 0.37 | 0.67 | 2.77 | 2.77 | 0 | 2.85 | 0.37 | 0.37
| Analytical results for Nb in OREAS 550b (Indicative Value 25.4 ppm) | [18 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 10 | 13 | 14 | 22 | 25 | 26 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | < 50 | < 50 | < 50 | 20.16 | 17.8 | < 70 | [11 empty] | 1
| 1 | 2 | < 50 | 60 | < 50 | 20.13 | 17.3 | < 70 | [11 empty] | 25
| 1 | 3 | < 50 | < 50 | 50 | 20.14 | 17.4 | < 70 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | 50 | 20.15 | 17.7 | < 70 | [11 empty] | 25.39
| 1 | 5 | < 50 | 50 | < 50 | 20.13 | 19.7 | < 70 | [11 empty] | 17
| 1 | 6 | < 50 | < 50 | < 50 | 20.15 | 18.6 | < 70
| Mean |  |  | 55 | 50 | 20.1433 | 18.0833
| Median |  |  | 55 | 50 | 20.145 | 17.75
| Std Dev. |  |  | 7.0711 | 0 | 0.012111 | 0.91524
| Rel.Std.Dev. |  |  | 12.86% | 0% | 0.06012% | 5.061%
| PDM3 |  |  | 116.6% | 96.93% | -20.66% | -28.78%
| Z-Score (Absolute) |  | 0.48 | 0.68 | 0.48 | 1.05 | 1.29 | 0.67
| Analytical results for Nd in OREAS 550b (Indicative Value 49.4 ppm) | [18 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 1
1
| 1 | 1 | 52.26 | 46.8 | < 170 | [14 empty] | 1
| 1 | 2 | 52.18 | 45.6 | < 170 | [14 empty] | 12
| 1 | 3 | 52.32 | 46.7 | < 170 | [14 empty] | 16
| 1 | 4 | 52.19 | 47.2 | < 170 | [14 empty] | 49.445
| 1 | 5 | 52.31 | 45.2 | < 170 | [14 empty] | 18
| 1 | 6 | 52.28 | 48.3 | < 170
| Mean |  | 52.2567 | 46.6333
| Median |  | 52.27 | 46.75
| Std Dev. |  | 0.059554 | 1.1183
| Rel.Std.Dev. |  | 0.114% | 2.398%
| PDM3 |  | 5.686% | -5.686%
| Z-Score (Absolute) |  | 0 | 0.67 | 4.01
| Analytical results for Ni in OREAS 550b (Certified Value 57 ppm) | [18 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 04 | 10 | 13 | 14 | 20 | 25 | 26 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 60 | < 50 | 60 | 90 | < 50 | 50 | 101 | < 80 | [9 empty] | 1
| 1 | 2 | 60 | < 50 | 50 | 300 | 70 | 60 | 98 | 80 | [9 empty] | 28
| 1 | 3 | 60 | < 50 | 50 | 230 | 60 | 60 | 95 | < 80 | [9 empty] | 16
| 1 | 4 | 60 | < 50 | < 50 | 110 | 60 | 60 | 93 | 80 | [9 empty] | 56.6781
| 1 | 5 | 50 | < 50 | < 50 | 150 | 80 | 70 | 95 | < 80 | [9 empty] | 82
| 1 | 6 | 60 | < 50 | 60 | 330 | 60 | 70 | 91 | < 80
| Mean |  | 58.3333 |  | 55 | 201.6667 | 66 | 61.6667 | 95.5 | 80
| Median |  | 60 |  | 55 | 190 | 60 | 60 | 95 | 80
| Std Dev. |  | 4.0825 |  | 5.7735 | 100.4822 | 8.9443 | 7.5277 | 3.5637 | 0
| Rel.Std.Dev. |  | 6.999% |  | 10.5% | 49.83% | 13.55% | 12.21% | 3.732% | 0%
| PDM3 |  | 2.92% |  | -2.961% | 255.8% | 16.45% | 8.802% | 68.5% | 41.15%
| Z-Score (Absolute) |  | 0.03 | 2.26 | 0.92 | 9.58 | 0.03 | 0.2 | 2.46 | 0.43
| Analytical results for P2O5 in OREAS 550b (Certified Value 0.115 wt.%) | [18 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.116 | 0.11 | 0.119 | 0.121 | 0.1169 | 0.1146 | 0.1 | 0.1 | 0.11 | 0.12 | 0.11051 | 0.104 | 0.119 | 0.11 | 0.17 | 0.13 |  | 1
| 1 | 2 | 0.117 | 0.12 | 0.12 | 0.118 | 0.1215 | 0.1146 | 0.1 | 0.11 | 0.11 | 0.12 | 0.11147 | 0.107 | 0.12 | 0.15 | 0.16 | 0.13 |  | #N/A
| 1 | 3 | 0.118 | 0.11 | 0.117 | 0.11 | 0.1192 | 0.1146 | 0.11 | 0.11 | 0.11 | 0.12 | 0.11196 | 0.107 | 0.12 | 0.13 | 0.16 | 0.13 |  | 16
| 1 | 4 | 0.117 | 0.11 | 0.117 | 0.13 | 0.1192 | 0.1146 | 0.1 | 0.11 | 0.11 | 0.12 | 0.11056 | 0.104 | 0.12 | 0.11 | 0.13 | 0.13 |  | 0.11534
| 1 | 5 | 0.119 | 0.12 | 0.117 | 0.12 | 0.1146 | 0.1146 | 0.1 | 0.11 | 0.11 | 0.12 | 0.11148 | 0.103 | 0.119 | 0.12 | 0.15 | 0.13 |  | 83
| 1 | 6 | 0.117 | 0.11 | 0.118 | 0.12 | 0.1192 | 0.1146 | 0.11 | 0.11 | 0.11 | 0.12 | 0.10953 | 0.112 | 0.121 | 0.13 | 0.17 | 0.13
| Mean |  | 0.11733 | 0.11333 | 0.118 | 0.11983 | 0.11843 | 0.1146 | 0.10333 | 0.10833 | 0.11 | 0.12 | 0.11092 | 0.10617 | 0.11983 | 0.125 | 0.15667 | 0.13
| Median |  | 0.117 | 0.11 | 0.1175 | 0.12 | 0.1192 | 0.1146 | 0.1 | 0.11 | 0.11 | 0.12 | 0.11102 | 0.1055 | 0.12 | 0.125 | 0.16 | 0.13
| Std Dev. |  | 0.0010328 | 0.005164 | 0.0012649 | 0.0064005 | 0.0023754 | 0 | 0.005164 | 0.0040825 | 0 | 0 | 0.00088617 | 0.0033116 | 0.00075277 | 0.015166 | 0.015055 | 0
| Rel.Std.Dev. |  | 0.8802% | 4.556% | 1.072% | 5.341% | 2.006% | 0% | 4.997% | 3.768% | 0% | 0% | 0.7989% | 3.119% | 0.6282% | 12.13% | 9.61% | 0%
| PDM3 |  | 1.73% | -1.738% | 2.308% | 3.898% | 2.684% | -0.6396% | -10.41% | -6.073% | -4.628% | 4.042% | -3.831% | -7.951% | 3.898% | 8.377% | 35.83% | 12.71%
| Z-Score (Absolute) |  | 0.04 | 0.53 | 0.04 | 0.26 | 0.09 | 0.38 | 1.74 | 1.14 | 0.93 | 0.28 | 0.82 | 1.4 | 0.26 | 0.89 | 4.75 | 1.5
| Analytical results for Pb in OREAS 550b (Indicative Value 48.2 ppm) | [18 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 23 | 25 | 26 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 1
1
| 1 | 1 | 100 | < 50 | < 50 | < 50 | < 100 | 12.9 | 190 | [10 empty] | 1
| 1 | 2 | 100 | < 50 | < 50 | < 50 | < 100 | 12.5 | < 90 | [10 empty] | 13
| 1 | 3 | 100 | < 50 | 50 | < 50 | < 100 | 13.2 | < 90 | [10 empty] | 16
| 1 | 4 | 100 | < 50 | 50 | < 50 | < 100 | 13.4 | < 90 | [10 empty] | 48.2458
| 1 | 5 | 100 | < 50 | < 50 | < 50 | < 100 | 13.6 | < 90 | [10 empty] | 19
| 1 | 6 | 100 | < 50 | < 50 | < 50 | < 100 | 13.8 | < 90
| Mean |  | 100 |  | 50 |  |  | 13.2333 | 190
| Median |  | 100 |  | 50 |  |  | 13.3 | 190
| Std Dev. |  | 0 |  | 0 |  |  | 0.4761
| Rel.Std.Dev. |  | 0% |  | 0% |  |  | 3.598%
| PDM3 |  | 107.3% |  | 3.636% |  |  | -72.57% | 293.8%
| Z-Score (Absolute) |  | 2.7 | 0.34 | 0 | 0.34 | 0.67 | 0.81 | 1.47
| Analytical results for Pr in OREAS 550b (Indicative Value 14.8 ppm) | [18 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 1
1
| 1 | 1 | 16.15 | 13.4 | < 170 | [14 empty] | 1
| 1 | 2 | 16.18 | 12.9 | < 170 | [14 empty] | 14
| 1 | 3 | 16.21 | 13.4 | < 170 | [14 empty] | 16
| 1 | 4 | 16.16 | 13.8 | < 170 | [14 empty] | 14.8308
| 1 | 5 | 16.19 | 13.3 | < 170 | [14 empty] | 20
| 1 | 6 | 16.18 | 14.1 | < 170
| Mean |  | 16.1783 | 13.4833
| Median |  | 16.18 | 13.4
| Std Dev. |  | 0.02137 | 0.41673
| Rel.Std.Dev. |  | 0.1321% | 3.091%
| PDM3 |  | 9.086% | -9.086%
| Z-Score (Absolute) |  | 0 | 0.67 | 16.66
| Analytical results for Rb in OREAS 550b (Indicative Value 60 ppm) | [18 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 13 | 14 | 25 | 26 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 60 | < 50 | 80.9 | < 180 | [13 empty] | 1
| 1 | 2 | 60 | < 50 | 80.3 | < 180 | [13 empty] | 15
| 1 | 3 | 90 | 60 | 82.6 | < 180 | [13 empty] | 16
| 1 | 4 | 90 | < 50 | 79.2 | < 180 | [13 empty] | 59.5333
| 1 | 5 | 60 | < 50 | 84.7 | < 180 | [13 empty] | 21
| 1 | 6 | 70 | < 50 | 83.9 | < 180
| Mean |  | 71.6667 | 60 | 81.9333
| Median |  | 65 | 60 | 81.75
| Std Dev. |  | 14.7196 |  | 2.1528
| Rel.Std.Dev. |  | 20.54% |  | 2.628%
| PDM3 |  | 20.38% | 0.7839% | 37.63%
| Z-Score (Absolute) |  | 0.35 | 3.13 | 0.35 | 1
| Analytical results for Re in OREAS 550b (Indicative Value < 0.1 ppm) | [18 empty] | Indicative
Re | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | < 0.1 | [16 empty] | 1
| 1 | 2 | < 0.1 | [16 empty] | 16
| 1 | 3 | < 0.1 | [16 empty] | 16
| 1 | 4 | < 0.1 | [16 empty] | < 0.1
| 1 | 5 | < 0.1 | [16 empty] | 22
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 550b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 25 | 26 | [11 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 10 | < 50 | < 50 | < 50 | 3 | < 80 | [11 empty] | 1
| 1 | 2 | < 10 | < 50 | < 50 | < 50 | 3 | < 80 | [11 empty] | 17
| 1 | 3 | < 10 | < 50 | < 50 | < 50 | 3 | < 80 | [11 empty] | 16
| 1 | 4 | < 10 | < 50 | < 50 | < 50 | 3 | 170 | [11 empty] | < 50
| 1 | 5 | < 10 | < 50 | < 50 | < 50 | 3 | < 80 | [11 empty] | 23
| 1 | 6 | < 10 | < 50 | < 50 | < 50 | 3 | < 80
| Mean | [5 empty] | 3 | 170
| Median | [5 empty] | 3 | 170
| Std Dev. | [5 empty] | 0
| Rel.Std.Dev. | [5 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 1.35 | 0 | 0 | 0 | 1.48 | 2.54
| Analytical results for Sc in OREAS 550b (Indicative Value 11.7 ppm) | [18 empty] | Indicative
Sc | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 1
1
| 1 | 1 | 12 | [16 empty] | 1
| 1 | 2 | 12 | [16 empty] | 18
| 1 | 3 | 14 | [16 empty] | 16
| 1 | 4 | 9 | [16 empty] | 11.6667
| 1 | 5 | 11 | [16 empty] | 24
| 1 | 6 | 12
| Mean |  | 11.6667
| Median |  | 12
| Std Dev. |  | 1.633
| Rel.Std.Dev. |  | 14%
| PDM3 |  | -2.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 550b (Indicative Value 3.5 ppm) | [18 empty] | Indicative
Se | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 3 | [16 empty] | 1
| 1 | 2 | 5 | [16 empty] | 22
| 1 | 3 | 4 | [16 empty] | 16
| 1 | 4 | 3 | [16 empty] | 3.5
| 1 | 5 | 3 | [16 empty] | 10
| 1 | 6 | 3
| Mean |  | 3.5
| Median |  | 3
| Std Dev. |  | 0.83666
| Rel.Std.Dev. |  | 23.9%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 550b (Certified Value 58.18 wt.%) | [18 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 58.23 | 57.6 | 58.65 | 57.689 | 58.59 | 57.074 | 58.1 | 56.2 | 58.7 | 58 | 57.6858 | 60.35 | 58.36 | 53.84 | 130.2099 | 55.39 | 58.3 | 1
| 1 | 2 | 58.25 | 58.01 | 58.97 | 57.716 | 59.55 | 57.48 | 57.9 | 58.5 | 58.6 | 58 | 57.8451 | 61.4 | 58.48 | 54.1 | 131.0634 | 53.93 | 57.5 | #N/A
| 1 | 3 | 58.23 | 58.42 | 58.16 | 57.75 | 59.58 | 57.93 | 58.2 | 58.6 | 58.9 | 58.4 | 57.8558 | 60.99 | 58.5 | 53.95 | 132.1865 | 57.43 | 57.7 | 16
| 1 | 4 | 58.08 | 57.88 | 58.35 | 57.71 | 59.56 | 57.48 | 57.7 | 57.7 | 59 | 58.3 | 57.2074 | 61.15 | 58.49 | 54.09 | 131.5351 | 55.13 | 57.9 | 58.1847
| 1 | 5 | 58.01 | 58.4 | 58.56 | 57.72 | 57.88 | 57.93 | 58.2 | 58.9 | 58.9 | 58 | 57.9749 | 59.96 | 58.4 | 53.94 | 131.4228 | 56.52 | 58 | 84
| 1 | 6 | 58.05 | 58.08 | 58.48 | 57.73 | 59.2 | 56.946 | 58.4 | 58.3 | 58.1 | 58.4 | 57.8099 | 60.89 | 58.47 | 54.1 | 131.8047 | 58.27 | 57.8
| Mean |  | 58.1417 | 58.065 | 58.5283 | 57.7192 | 59.06 | 57.4733 | 58.0833 | 58.0333 | 58.7 | 58.1833 | 57.7298 | 60.79 | 58.45 | 54.0033 | 131.3704 | 56.1117 | 57.8667
| Median |  | 58.155 | 58.045 | 58.52 | 57.718 | 59.375 | 57.48 | 58.15 | 58.4 | 58.8 | 58.15 | 57.8275 | 60.94 | 58.475 | 54.02 | 131.479 | 55.955 | 57.85
| Std Dev. |  | 0.10666 | 0.31367 | 0.27593 | 0.020361 | 0.69187 | 0.41346 | 0.24833 | 0.98319 | 0.32863 | 0.20412 | 0.27218 | 0.53557 | 0.056569 | 0.1093 | 0.68198 | 1.6005 | 0.27325
| Rel.Std.Dev. |  | 0.1835% | 0.5402% | 0.4714% | 0.03528% | 1.171% | 0.7194% | 0.4275% | 1.694% | 0.5599% | 0.3508% | 0.4715% | 0.881% | 0.09678% | 0.2024% | 0.5191% | 2.852% | 0.4722%
| PDM3 |  | -0.07388% | -0.2056% | 0.5907% | -0.8% | 1.504% | -1.223% | -0.1741% | -0.2601% | 0.8857% | -0.002266% | -0.7817% | 4.478% | 0.456% | -7.186% | 125.8% | -3.563% | -0.5465%
| Z-Score (Absolute) |  | 0.11 | 0.03 | 0.82 | 0.67 | 1.8 | 1.12 | 0 | 0.09 | 1.13 | 0.18 | 0.65 | 4.98 | 0.67 | 7.5 | 134.78 | 3.63 | 0.4
| Analytical results for Sm in OREAS 550b (Indicative Value 7.05 ppm) | [18 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 6.7 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 6.9 | < 170 | [14 empty] | 5
| 1 | 3 | < 10 | 7.1 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 6.9 | < 170 | [14 empty] | 7.05
| 1 | 5 | < 10 | 7 | < 170 | [14 empty] | 11
| 1 | 6 | < 10 | 7.7 | < 170
| Mean |  |  | 7.05
| Median |  |  | 6.95
| Std Dev. |  |  | 0.34496
| Rel.Std.Dev. |  |  | 4.893%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0 | 26.05
| Analytical results for Sn in OREAS 550b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 10 | 13 | 14 | 23 | 25 | 26 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 50 | 3050 | < 50 | < 100 | 2.5 | < 80 | [11 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 100 | 2.4 | < 80 | [11 empty] | 6
| 1 | 3 | < 50 | < 50 | < 50 | < 100 | 2.5 | < 80 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 100 | 2.5 | < 80 | [11 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | < 100 | 2.6 | < 80 | [11 empty] | 12
| 1 | 6 | < 50 | < 50 | < 50 | < 100 | 2.5 | < 80
| Mean |  |  | 3050 |  |  | 2.5
| Median |  |  | 3050 |  |  | 2.5
| Std Dev. | [5 empty] | 0.063246
| Rel.Std.Dev. | [5 empty] | 2.53%
PDM3
| Z-Score (Absolute) |  | 0.39 | 26.81 | 0.39 | 0.96 | 1.6 | 0.39
| Analytical results for SO3 in OREAS 550b (Certified Value 3 wt.%) | [18 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 04 | 07 | 10 | 11 | 13 | 14 | 19 | 22 | 26 | [7 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [7 empty] | 2
3
| 1 | 1 | 2.83 | 2.9941 | 2.9217 | 2.997 | 2.91 | 2.947 | 2.997 | 3.156 | 2.8 | 3.1 | [7 empty] | 1
| 1 | 2 | 2.64 | 3.114 | 2.8468 | 2.997 | 3.03 | 2.947 | 2.997 | 3.196 | 2.56 | 3 | [7 empty] | #N/A
| 1 | 3 | 2.74 | 3.069 | 2.8967 | 3.047 | 3.03 | 2.947 | 3.022 | 3.179 | 2.66 | 3 | [7 empty] | 16
| 1 | 4 | 2.83 | 2.9891 | 2.8967 | 3.047 | 2.99 | 2.972 | 3.047 | 3.184 | 2.59 | 3.1 | [7 empty] | 3.0023
| 1 | 5 | 2.94 | 3.1365 | 2.9467 | 2.997 | 3.05 | 2.972 | 2.997 | 3.118 | 2.71 | 3 | [7 empty] | 85
| 1 | 6 | 2.89 | 3.2713 | 2.8468 | 2.997 | 3.04 | 2.897 | 3.022 | 3.196 | 2.53 | 3
| Mean |  | 2.8117 | 3.0957 | 2.8926 | 3.0137 | 3.0083 | 2.947 | 3.0137 | 3.1715 | 2.6417 | 3.0333
| Median |  | 2.83 | 3.0915 | 2.8967 | 2.997 | 3.03 | 2.947 | 3.0095 | 3.1815 | 2.625 | 3
| Std Dev. |  | 0.10759 | 0.1051 | 0.040006 | 0.02582 | 0.052313 | 0.027386 | 0.020412 | 0.030052 | 0.10187 | 0.05164
| Rel.Std.Dev. |  | 3.827% | 3.395% | 1.383% | 0.8568% | 1.739% | 0.9293% | 0.6773% | 0.9476% | 3.856% | 1.702%
| PDM3 |  | -6.349% | 3.11% | -3.655% | 0.3787% | 0.2011% | -1.842% | 0.3787% | 5.636% | -12.01% | 1.034%
| Z-Score (Absolute) |  | 1.8 | 0.77 | 1.07 | 0.02 | 0.02 | 0.58 | 0.02 | 1.45 | 3.34 | 0.2
| Analytical results for SrO in OREAS 550b (Certified Value 102 ppm) | [18 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 10 | 11 | 13 | 14 | 15 | 19 | 22 | 25 | 26 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 120 | 100 | 120 | 120 | < 50 | 90 | 100 | 81.6 | < 100 | [8 empty] | 1
| 1 | 2 | < 120 | 100 | 120 | 120 | < 50 | 100 | 200 | 81.6 | < 100 | [8 empty] | #N/A
| 1 | 3 | < 120 | 100 | 120 | 120 | < 50 | 90 | 100 | 81.6 | < 100 | [8 empty] | 16
| 1 | 4 | < 120 | 100 | 120 | 120 | < 50 | 100 | 100 | 81.6 | < 100 | [8 empty] | 102.1716
| 1 | 5 | < 120 | 100 | 120 | 120 | < 50 | 90 | 100 | 87.5 | < 100 | [8 empty] | 86
| 1 | 6 | < 120 | 100 | 120 | 120 | < 50 | 90 | 100 | 85.1 | < 100
| Mean |  |  | 100 | 120 | 120 |  | 93.3333 | 116.6667 | 83.1667
| Median |  |  | 100 | 120 | 120 |  | 90 | 100 | 81.6
| Std Dev. |  |  | 0 | 0 | 0 |  | 5.164 | 40.8248 | 2.543
| Rel.Std.Dev. |  |  | 0% | 0% | 0% |  | 5.533% | 34.99% | 3.058%
| PDM3 |  |  | -2.125% | 17.45% | 17.45% |  | -8.65% | 14.19% | -18.6%
| Z-Score (Absolute) |  | 0.93 | 0.18 | 0.67 | 0.67 | 1.85 | 0 | 0.63 | 0.27 | 1.17
| Analytical results for Ta in OREAS 550b (Indicative Value 34.8 ppm) | [18 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 10 | 25 | 26 | [14 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 1
1
| 1 | 1 | < 20 | 1.2 | 80 | [14 empty] | 1
| 1 | 2 | < 20 | 1.3 | < 80 | [14 empty] | 7
| 1 | 3 | < 20 | 1.3 | 80 | [14 empty] | 16
| 1 | 4 | < 20 | 1.4 | 80 | [14 empty] | 34.7972
| 1 | 5 | < 20 | 1.5 | < 80 | [14 empty] | 13
| 1 | 6 | < 20 | 1.4 | 80
| Mean |  |  | 1.35 | 80
| Median |  |  | 1.35 | 80
| Std Dev. |  |  | 0.10488 | 0
| Rel.Std.Dev. |  |  | 7.769% | 0%
| PDM3 |  |  | -96.12% | 129.9%
| Z-Score (Absolute) |  | 0 | 0.67 | 4.54
| Analytical results for Tb in OREAS 550b (Indicative Value 0.9 ppm) | [18 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 0.86 | [15 empty] | 1
| 1 | 2 | < 10 | 0.84 | [15 empty] | 8
| 1 | 3 | < 10 | 0.94 | [15 empty] | 16
| 1 | 4 | < 10 | 0.91 | [15 empty] | 0.90167
| 1 | 5 | < 10 | 0.89 | [15 empty] | 14
| 1 | 6 | < 10 | 0.97
| Mean |  |  | 0.90167
| Median |  |  | 0.9
| Std Dev. |  |  | 0.048751
| Rel.Std.Dev. |  |  | 5.407%
| PDM3 |  |  | -4.441e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 550b (Indicative Value < 0.1 ppm) | [18 empty] | Indicative
Te | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | < 0.1 | [16 empty] | 1
| 1 | 2 | < 0.1 | [16 empty] | 9
| 1 | 3 | < 0.1 | [16 empty] | 16
| 1 | 4 | < 0.1 | [16 empty] | < 0.1
| 1 | 5 | < 0.1 | [16 empty] | 15
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 550b (Indicative Value 13.8 ppm) | [18 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 10 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [13 empty] | 1
1
| 1 | 1 | < 20 | 14.56 | 12.9 | 260 | [13 empty] | 1
| 1 | 2 | < 20 | 14.39 | 12.8 | 180 | [13 empty] | 10
| 1 | 3 | < 20 | 14.49 | 13.2 | 180 | [13 empty] | 16
| 1 | 4 | < 20 | 14.51 | 13 | 180 | [13 empty] | 13.8117
| 1 | 5 | < 20 | 14.52 | 13.4 | 90 | [13 empty] | 16
| 1 | 6 | < 20 | 14.57 | 13.4 | < 90
| Mean |  |  | 14.5067 | 13.1167 | 178
| Median |  |  | 14.515 | 13.1 | 180
| Std Dev. |  |  | 0.064704 | 0.25626 | 60.1664
| Rel.Std.Dev. |  |  | 0.446% | 1.954% | 33.8%
| PDM3 |  |  | 5.032% | -5.032% | 1189%
| Z-Score (Absolute) |  | 1.14 | 0.21 | 0.21 | 41.89
| Analytical results for TiO2 in OREAS 550b (Certified Value 0.646 wt.%) | [18 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.64 | 0.64 | 0.59 | 0.663 | 0.65 | 0.6506 | 0.64 | 0.6 | 0.63 | 0.65 | 0.63695 | 0.65 | 0.66 | 0.66 | 0.539 | 0.65 |  | 1
| 1 | 2 | 0.64 | 0.64 | 0.6 | 0.656 | 0.66 | 0.6672 | 0.62 | 0.63 | 0.62 | 0.64 | 0.64266 | 0.65 | 0.69 | 0.67 | 0.562 | 0.63 |  | #N/A
| 1 | 3 | 0.64 | 0.65 | 0.59 | 0.65 | 0.65 | 0.6506 | 0.62 | 0.63 | 0.63 | 0.65 | 0.63706 | 0.66 | 0.67 | 0.67 | 0.593 | 0.63 |  | 16
| 1 | 4 | 0.64 | 0.64 | 0.59 | 0.66 | 0.65 | 0.6506 | 0.63 | 0.62 | 0.64 | 0.65 | 0.63153 | 0.65 | 0.67 | 0.68 | 0.558 | 0.64 |  | 0.64592
| 1 | 5 | 0.64 | 0.66 | 0.6 | 0.64 | 0.63 | 0.6672 | 0.62 | 0.63 | 0.63 | 0.65 | 0.64457 | 0.64 | 0.67 | 0.68 | 0.584 | 0.64 |  | 87
| 1 | 6 | 0.63 | 0.65 | 0.61 | 0.65 | 0.65 | 0.6672 | 0.62 | 0.64 | 0.63 | 0.65 | 0.63207 | 0.65 | 0.66 | 0.66 | 0.608 | 0.63
| Mean |  | 0.63833 | 0.64667 | 0.59667 | 0.65317 | 0.64833 | 0.6589 | 0.625 | 0.625 | 0.63 | 0.64833 | 0.63747 | 0.65 | 0.67 | 0.67 | 0.574 | 0.63667
| Median |  | 0.64 | 0.645 | 0.595 | 0.653 | 0.65 | 0.6589 | 0.62 | 0.63 | 0.63 | 0.65 | 0.63701 | 0.65 | 0.67 | 0.67 | 0.573 | 0.635
| Std Dev. |  | 0.0040825 | 0.008165 | 0.008165 | 0.0083046 | 0.0098319 | 0.0090922 | 0.0083666 | 0.013784 | 0.0063246 | 0.0040825 | 0.0053315 | 0.0063246 | 0.010954 | 0.0089443 | 0.025464 | 0.008165
| Rel.Std.Dev. |  | 0.6396% | 1.263% | 1.368% | 1.271% | 1.516% | 1.38% | 1.339% | 2.205% | 1.004% | 0.6297% | 0.8364% | 0.973% | 1.635% | 1.335% | 4.436% | 1.282%
| PDM3 |  | -1.174% | 0.1157% | -7.625% | 1.122% | 0.3737% | 2.01% | -3.239% | -3.239% | -2.465% | 0.3737% | -1.308% | 0.6317% | 3.728% | 3.728% | -11.13% | -1.433%
| Z-Score (Absolute) |  | 0.24 | 0.24 | 2.67 | 0.62 | 0.34 | 0.95 | 1.02 | 1.02 | 0.73 | 0.34 | 0.29 | 0.44 | 1.6 | 1.6 | 3.99 | 0.34
| Analytical results for Tl in OREAS 550b (Indicative Value 0.47 ppm) | [18 empty] | Indicative
Tl | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 0.46 | [16 empty] | 1
| 1 | 2 | 0.45 | [16 empty] | 26
| 1 | 3 | 0.43 | [16 empty] | 16
| 1 | 4 | 0.48 | [16 empty] | 0.465
| 1 | 5 | 0.49 | [16 empty] | 17
| 1 | 6 | 0.48
| Mean |  | 0.465
| Median |  | 0.47
| Std Dev. |  | 0.022583
| Rel.Std.Dev. |  | 4.857%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 550b (Indicative Value 0.44 ppm) | [18 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 0.4 | [15 empty] | 1
| 1 | 2 | < 10 | 0.42 | [15 empty] | 12
| 1 | 3 | < 10 | 0.44 | [15 empty] | 16
| 1 | 4 | < 10 | 0.45 | [15 empty] | 0.44333
| 1 | 5 | < 10 | 0.45 | [15 empty] | 18
| 1 | 6 | < 10 | 0.5
| Mean |  |  | 0.44333
| Median |  |  | 0.445
| Std Dev. |  |  | 0.033862
| Rel.Std.Dev. |  |  | 7.638%
| PDM3 |  |  | 6.661e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 550b (Indicative Value 7.2 ppm) | [18 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 10 | 25 | 26 | [14 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | 10 | 8.4 | 170 | [14 empty] | 1
| 1 | 2 | < 10 | 8.43 | 80 | [14 empty] | 13
| 1 | 3 | < 10 | 8.49 | < 80 | [14 empty] | 16
| 1 | 4 | < 10 | 8.6 | 170 | [14 empty] | 7.2033
| 1 | 5 | < 10 | 8.69 | < 80 | [14 empty] | 19
| 1 | 6 | < 10 | 8.83 | < 80
| Mean |  | 10 | 8.5733 | 140
| Median |  | 10 | 8.545 | 170
| Std Dev. |  |  | 0.16597 | 51.9615
| Rel.Std.Dev. |  |  | 1.936% | 37.12%
| PDM3 |  | 38.82% | 19.02% | 1844%
| Z-Score (Absolute) |  | 0.67 | 0 | 20.5
| Analytical results for V2O5 in OREAS 550b (Certified Value 363 ppm) | [18 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 06 | 10 | 13 | 14 | 19 | 25 | 26 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 310 | 300 | 360 | 360 | 180 | 400 | 348.1 | 500 | [9 empty] | 1
| 1 | 2 | 320 | 300 | 360 | 360 | 180 | 460 | 357 | 400 | [9 empty] | #N/A
| 1 | 3 | 290 | 300 | 360 | 360 | 180 | 420 | 360.6 | 500 | [9 empty] | 16
| 1 | 4 | 290 | 300 | 180 | 360 | 180 | 460 | 358.8 | 400 | [9 empty] | 362.5762
| 1 | 5 | 320 | 300 | 360 | 360 | 180 | 440 | 348.1 | 500 | [9 empty] | 88
| 1 | 6 | 310 | 300 | 360 | 360 | 180 | 360 | 369.5 | 500
| Mean |  | 306.6667 | 300 | 330 | 360 | 180 | 423.3333 | 357.0167 | 466.6667
| Median |  | 310 | 300 | 360 | 360 | 180 | 430 | 357.9 | 500
| Std Dev. |  | 13.6626 | 0 | 73.4847 | 0 | 0 | 38.8158 | 8.1347 | 51.6398
| Rel.Std.Dev. |  | 4.455% | 0% | 22.27% | 0% | 0% | 9.169% | 2.279% | 11.07%
| PDM3 |  | -15.42% | -17.26% | -8.985% | -0.7105% | -50.36% | 16.76% | -1.533% | 28.71%
| Z-Score (Absolute) |  | 0.62 | 0.73 | 0.26 | 0.26 | 2.84 | 1.41 | 0.26 | 2.16
| Analytical results for W in OREAS 550b (Indicative Value < 10 ppm) | [18 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 10 | 13 | 14 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | 8 | < 10 | < 10 | 3.1 | < 80 | [12 empty] | 1
| 1 | 2 | < 8 | < 10 | < 10 | 2.9 | < 80 | [12 empty] | 14
| 1 | 3 | < 8 | < 10 | < 10 | 2.9 | < 80 | [12 empty] | 16
| 1 | 4 | < 8 | < 10 | < 10 | 3 | < 80 | [12 empty] | < 10
| 1 | 5 | < 8 | < 10 | < 10 | 3.6 | < 80 | [12 empty] | 20
| 1 | 6 | < 8 | < 10 | < 10 | 3.2 | < 80
| Mean |  | 8 |  |  | 3.1167
| Median |  | 8 |  |  | 3.05
| Std Dev. |  |  |  |  | 0.26394
| Rel.Std.Dev. |  |  |  |  | 8.469%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 3.4 | 62.47
| Analytical results for Y in OREAS 550b (Indicative Value 38.9 ppm) | [18 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | 39 | 47 | < 10 | 27.5 | < 160 | [12 empty] | 1
| 1 | 2 | < 39 | 55 | < 10 | 28.2 | < 160 | [12 empty] | 15
| 1 | 3 | 47 | 63 | < 10 | 29.1 | < 160 | [12 empty] | 16
| 1 | 4 | 39 | 55 | < 10 | 28.6 | < 160 | [12 empty] | 38.8834
| 1 | 5 | < 39 | 47 | < 10 | 29.1 | < 160 | [12 empty] | 21
| 1 | 6 | 39 | 55 | < 10 | 29.8 | < 160
| Mean |  | 41 | 53.6667 |  | 28.7167
| Median |  | 39 | 55 |  | 28.85
| Std Dev. |  | 4 | 6.0222 |  | 0.80353
| Rel.Std.Dev. |  | 9.756% | 11.22% |  | 2.798%
| PDM3 |  | 5.444% | 38.02% |  | -26.15%
| Z-Score (Absolute) |  | 0 | 0.67 | 1 | 0.19 | 1.53
| Analytical results for Yb in OREAS 550b (Indicative Value 2.85 ppm) | [18 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 2.9 | [15 empty] | 1
| 1 | 2 | < 10 | 2.7 | [15 empty] | 16
| 1 | 3 | < 10 | 2.9 | [15 empty] | 16
| 1 | 4 | < 10 | 2.8 | [15 empty] | 2.85
| 1 | 5 | < 10 | 2.8 | [15 empty] | 22
| 1 | 6 | < 10 | 3
| Mean |  |  | 2.85
| Median |  |  | 2.85
| Std Dev. |  |  | 0.10488
| Rel.Std.Dev. |  |  | 3.68%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 550b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 07 | 10 | 13 | 14 | 20 | 23 | 25 | 26 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | 60 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 26 | < 80 | [8 empty] | 1
| 1 | 2 | 60 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | 23 | < 80 | [8 empty] | 17
| 1 | 3 | 50 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 23 | < 80 | [8 empty] | 16
| 1 | 4 | 60 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 23 | < 80 | [8 empty] | < 50
| 1 | 5 | 60 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | 26 | < 80 | [8 empty] | 23
| 1 | 6 | 60 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | 26 | < 80
| Mean |  | 58.3333 |  |  |  |  | 20 |  | 24.5
| Median |  | 60 |  |  |  |  | 20 |  | 24.5
| Std Dev. |  | 4.0825 |  |  |  |  | 0 |  | 1.6432
| Rel.Std.Dev. |  | 6.999% |  |  |  |  | 0% |  | 6.707%
PDM3
| Z-Score (Absolute) |  | 44.95 | 0 | 0 | 0 | 0 | 16.86 | 33.72 | 0.67 | 20.45
| Analytical results for Zr in OREAS 550b (Certified Value 184 ppm) | [18 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 10 | 13 | 14 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 150 | 200 | 200 | 217 | < 150 | [12 empty] | 1
| 1 | 2 | 150 | 200 | 100 | 217 | < 150 | [12 empty] | 22
| 1 | 3 | 150 | 200 | 200 | 216 | < 150 | [12 empty] | 16
| 1 | 4 | 150 | 200 | 200 | 210 | < 150 | [12 empty] | 184.2244
| 1 | 5 | 150 | 200 | 200 | 187 | < 150 | [12 empty] | 89
| 1 | 6 | 150 | 200 | 200 | 186 | < 150
| Mean |  | 150 | 200 | 183.3333 | 205.5
| Median |  | 150 | 200 | 200 | 213
| Std Dev. |  | 0 | 0 | 40.8248 | 14.9499
| Rel.Std.Dev. |  | 0% | 0% | 22.27% | 7.275%
| PDM3 |  | -18.58% | 8.563% | -0.4837% | 11.55%
| Z-Score (Absolute) |  | 1.07 | 0.51 | 0 | 0.67 | 3.32
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 550b (Certified Value 9.97 wt.%) | [22 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | LOI*TGA | BF*OES | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | 2
3
| 1 | 1 | 9.55 | 9.43 | 10.71 | 11.21 | 10.5635 | 9.9095 | 10.15 | 9.7 | 10.9065 | 10.13 | 9.61 | 8.95 | 9.45 | 9.2667 | 9.72 | 9.93 | 10.2 | 10.18 | 10.35 | 9.843 | 9.56 | 1
| 1 | 2 | 9.59 | 9.55 | 10.72 | 11.03 | 10.5489 | 10.1413 | 10.12 | 9.8 | 10.9641 | 10.14 | 9.61 | 9.17 | 9.37 | 9.26 | 9.77 | 10.1 | 10.2 | 10.17 | 10.19 | 9.562 | 9.55 | #N/A
| 1 | 3 | 9.6 | 9.45 | 10.72 | 10.75 | 10.6417 | 10.2357 | 10.16 | 9.7 | 10.9695 | 10.21 | 9.35 | 9.14 | 9.33 | 9.2505 | 9.86 | 10.36 | 10.2 | 10.2 | 9.988 | 9.648 | 9.58 | 16
| 1 | 4 | 9.49 | 9.42 | 10.84 | 11.27 | 10.7774 | 10.1888 | 10.33 | 9.7 | 10.9719 | 10.15 | 9.71 | 9.15 | 9.42 | 9.2314 | 9.81 | 10.42 | 10.1 | 10.04 | 9.96 | 9.87 | 9.6 | 9.9748
| 1 | 5 | 9.56 | 9.55 | 10.47 | 11.26 | 11.0425 | 10.0019 | 10.31 | 9.8 | 10.8906 | 10.15 | 9.62 | 9.16 | 9.39 | 9.2567 | 9.77 | 10.47 | 10.2 | 10.24 | 10.1 | 8.742 | 9.6 | 91
| 1 | 6 | 9.58 | 9.3 | 10.49 | 11.15 | 11.55 | 9.9324 | 10.37 | 9.7 | 10.8991 | 9.85 | 9.69 | 9.15 | 9.4 | 9.2894 | 9.74 | 10.55 | 10.2 | 10 | 10 | 8.985 | 9.58
| Mean |  | 9.5617 | 9.45 | 10.6583 | 11.1117 | 10.854 | 10.0683 | 10.24 | 9.7333 | 10.9336 | 10.105 | 9.5983 | 9.12 | 9.3933 | 9.2591 | 9.7783 | 10.305 | 10.1833 | 10.1383 | 10.098 | 9.4417 | 9.5783
| Median |  | 9.57 | 9.44 | 10.715 | 11.18 | 10.7095 | 10.0716 | 10.235 | 9.7 | 10.9353 | 10.145 | 9.615 | 9.15 | 9.395 | 9.2584 | 9.77 | 10.39 | 10.2 | 10.175 | 10.05 | 9.605 | 9.58
| Std Dev. |  | 0.039707 | 0.093595 | 0.14634 | 0.19783 | 0.38687 | 0.13853 | 0.10844 | 0.05164 | 0.038646 | 0.12802 | 0.12906 | 0.083905 | 0.041312 | 0.019117 | 0.050365 | 0.23923 | 0.040825 | 0.095586 | 0.15015 | 0.46897 | 0.020412
| Rel.Std.Dev. |  | 0.4153% | 0.9904% | 1.373% | 1.78% | 3.564% | 1.376% | 1.059% | 0.5305% | 0.3535% | 1.267% | 1.345% | 0.92% | 0.4398% | 0.2065% | 0.5151% | 2.321% | 0.4009% | 0.9428% | 1.487% | 4.967% | 0.2131%
| PDM3 |  | -4.142% | -5.261% | 6.853% | 11.4% | 8.814% | 0.9373% | 2.659% | -2.421% | 9.613% | 1.305% | -3.774% | -8.569% | -5.829% | -7.175% | -1.969% | 3.311% | 2.091% | 1.64% | 1.235% | -5.345% | -3.975%
| Z-Score (Absolute) |  | 0.7 | 0.85 | 0.81 | 1.44 | 1.08 | 0 | 0.24 | 0.46 | 1.19 | 0.05 | 0.65 | 1.31 | 0.93 | 1.11 | 0.4 | 0.33 | 0.16 | 0.1 | 0.04 | 0.86 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 550b (Indicative Value 0.633 ppm) | [23 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 17 | 22 | 25 | 26 | [18 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*MS | [18 empty] | 3
3
| 1 | 1 | < 1 | 1.07 | 0.2 | < 10 | [18 empty] | 1
| 1 | 2 | < 1 | 1.09 | 0.1 | < 10 | [18 empty] | 4
| 1 | 3 | < 1 | 1.07 | 0.1 | < 10 | [18 empty] | 16
| 1 | 4 | < 1 | 1.09 | 0.2 | < 10 | [18 empty] | 0.6325
| 1 | 5 | < 1 | 1.08 | 0.2 | < 10 | [18 empty] | 26
| 1 | 6 | < 1 | 1.09 | 0.3 | < 10
| Mean |  |  | 1.0817 | 0.18333
| Median |  |  | 1.085 | 0.2
| Std Dev. |  |  | 0.0098319 | 0.075277
| Rel.Std.Dev. |  |  | 0.909% | 41.06%
| PDM3 |  |  | 71.01% | -71.01%
| Z-Score (Absolute) |  | 0.44 | 0.44 | 0.91 | 6.32
| Analytical results for Al in OREAS 550b (Certified Value 5.88 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 5.94 | 6.123 | 6.15 | 5.8 | 5.8166 | 5.71 | 5.95 | 5.67 | 5.7023 | 5.822 | 5.716 | 6.139 | 5.875 | 6.007 | 5.945 | 5.76 | 5.9 | 6.0206 | 5.92 | 0.91 | 5.83 | 5.91 | 1
| 1 | 2 | 6.09 | 6.07 | 6.01 | 5.91 | 5.8159 | 5.57 | 5.94 | 5.62 | 5.7107 | 5.795 | 5.98 | 6.192 | 5.927 | 6.007 | 5.9675 | 5.98 | 5.88 | 6.0045 | 5.81 | 0.93 | 5.9 | 5.85 | #N/A
| 1 | 3 | 5.93 | 6.065 | 6.03 | 5.84 | 5.8165 | 5.72 | 5.96 | 5.74 | 5.7661 | 5.848 | 5.901 | 6.219 | 5.927 | 5.954 | 5.8794 | 5.87 | 5.82 | 6.1697 | 5.88 | 0.95 | 5.75 | 5.79 | 16
| 1 | 4 | 5.83 | 6.245 | 5.99 | 5.74 | 5.817 | 6.15 | 5.93 | 5.63 | 5.6199 | 5.901 | 5.875 | 6.113 | 5.822 | 6.007 | 6.0308 | 5.93 | 5.83 | 6.2194 | 5.74 | 0.92 | 5.52 | 5.78 | 5.8847
| 1 | 5 | 5.96 | 6.16 | 5.96 | 5.87 | 5.8104 | 5.55 | 5.94 | 5.72 | 5.7053 | 5.901 | 5.875 | 6.139 | 5.875 | 6.139 | 5.9029 | 5.92 | 6.18 | 6.0965 | 5.78 | 0.94 | 5.43 | 5.73 | 93
| 1 | 6 | 5.92 | 6.134 | 5.99 | 5.88 | 5.8175 | 6.06 | 5.97 | 5.69 | 5.6158 | 5.901 | 5.848 | 6.272 | 5.875 | 6.06 | 5.8733 | 5.84 | 6.02 | 6.1955 | 5.89 | 9.93 | 5.47 | 5.94
| Mean |  | 5.945 | 6.1328 | 6.0217 | 5.84 | 5.8156 | 5.7933 | 5.9483 | 5.6783 | 5.6867 | 5.8613 | 5.8658 | 6.179 | 5.8835 | 6.029 | 5.9331 | 5.8833 | 5.9383 | 6.1177 | 5.8367 | 2.43 | 5.65 | 5.8333
| Median |  | 5.935 | 6.1285 | 6 | 5.855 | 5.8165 | 5.715 | 5.945 | 5.68 | 5.7038 | 5.8745 | 5.875 | 6.1655 | 5.875 | 6.007 | 5.9239 | 5.895 | 5.89 | 6.1331 | 5.845 | 0.935 | 5.635 | 5.82
| Std Dev. |  | 0.084083 | 0.066282 | 0.067057 | 0.061644 | 0.0026395 | 0.25288 | 0.01472 | 0.047924 | 0.058229 | 0.046573 | 0.086263 | 0.059957 | 0.039455 | 0.063463 | 0.060418 | 0.077632 | 0.13834 | 0.091424 | 0.070616 | 3.6743 | 0.2013 | 0.081158
| Rel.Std.Dev. |  | 1.414% | 1.081% | 1.114% | 1.056% | 0.04539% | 4.365% | 0.2475% | 0.844% | 1.024% | 0.7946% | 1.471% | 0.9703% | 0.6706% | 1.053% | 1.018% | 1.32% | 2.33% | 1.494% | 1.21% | 151.2% | 3.563% | 1.391%
| PDM3 |  | 1.025% | 4.217% | 2.328% | -0.7595% | -1.173% | -1.553% | 1.081% | -3.507% | -3.365% | -0.397% | -0.3205% | 5.001% | -0.02031% | 2.452% | 0.8233% | -0.02315% | 0.9115% | 3.96% | -0.8162% | -58.71% | -3.988% | -0.8728%
| Z-Score (Absolute) |  | 0.66 | 2.42 | 1.38 | 0.32 | 0.55 | 0.76 | 0.69 | 1.83 | 1.76 | 0.12 | 0.08 | 2.85 | 0.08 | 1.44 | 0.55 | 0.08 | 0.6 | 2.27 | 0.35 | 32.2 | 2.1 | 0.39
| Analytical results for As in OREAS 550b (Certified Value 60 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 22 | 25 | 26 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS |  |  |  |  | 0
1
| 1 | 1 | 60 | < 100 | 58 | 60 | 62 | 70 | 57.5846 | 100 | < 100 | 100 | 100 | < 100 | 53 | 59 | < 100 | 42.2 | 55 | 61 |  |  |  |  | 1
| 1 | 2 | 64 | < 100 | 58 | 60 | 61 | 70 | 59.9906 | 100 | < 100 | 100 | 200 | < 100 | 53 | 56 | < 100 | 42.3 | 51 | 63 |  |  |  |  | 5
| 1 | 3 | 60 | < 100 | 64 | 60 | 61 | 70 | 58.2711 | 100 | < 100 | 100 | 100 | < 100 | 51 | 60 | < 100 | 42.2 | 56 | 64 |  |  |  |  | 16
| 1 | 4 | 62 | < 100 | 63 | 60 | 61 | 60 | 60.0966 | 100 | 100 | 100 | 100 | < 100 | 53 | 57 | < 100 | 42.2 | 59 | 58 |  |  |  |  | 59.9998
| 1 | 5 | 62 | < 100 | 59 | 60 | 62 | 70 | 58.0599 | 100 | < 100 | 100 | 100 | < 100 | 51 | 59 | < 100 | 42.2 | 61 | 67 |  |  |  |  | 94
| 1 | 6 | 61 | < 100 | 69 | 60 | 62 | 60 | 57.9853 | 100 | < 100 | 100 | 100 | < 100 | 50 | 64 | < 100 | 42.3 | 56 | 62
| Mean |  | 61.5 |  | 61.8333 | 60 | 61.5 | 66.6667 | 58.6647 | 100 | 100 | 100 | 116.6667 |  | 51.8333 | 59.1667 |  | 42.2333 | 56.3333 | 62.5
| Median |  | 61.5 |  | 61 | 60 | 61.5 | 70 | 58.1655 | 100 | 100 | 100 | 100 |  | 52 | 59 |  | 42.2 | 56 | 62.5
| Std Dev. |  | 1.5166 |  | 4.3551 | 0 | 0.54772 | 5.164 | 1.0915 | 0 |  | 0 | 40.8248 |  | 1.3292 | 2.7869 |  | 0.05164 | 3.4448 | 3.0166
| Rel.Std.Dev. |  | 2.466% |  | 7.043% | 0% | 0.8906% | 7.746% | 1.861% | 0% |  | 0% | 34.99% |  | 2.564% | 4.71% |  | 0.1223% | 6.115% | 4.827%
| PDM3 |  | 2.5% |  | 3.056% | 0.0003315% | 2.5% | 11.11% | -2.225% | 66.67% | 66.67% | 66.67% | 94.45% |  | -13.61% | -1.389% |  | -29.61% | -6.111% | 4.167%
| Z-Score (Absolute) |  | 0.25 | 1.25 | 0.29 | 0.05 | 0.25 | 0.92 | 0.12 | 5.27 | 0.16 | 5.27 | 7.45 | 1.25 | 1.01 | 0.05 | 1.25 | 2.26 | 0.42 | 0.38
| Analytical results for B in OREAS 550b (Certified Value 171 ppm) | [23 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 08 | 09 | 19 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | [15 empty] | 0
0
| 1 | 1 | 173 | 180 | 180 | 157.447 | 207 | 43 | 142 | [15 empty] | 1
| 1 | 2 | 180 | 170 | 170 | 156.1399 | 204 | 40 | 135 | [15 empty] | 1
| 1 | 3 | 180 | 170 | 180 | 159.5636 | 201 | 33 | 143 | [15 empty] | 16
| 1 | 4 | 183 | 170 | 180 | 156.4427 | 201 | 33 | 172 | [15 empty] | 170.7019
| 1 | 5 | 175 | 170 | 180 | 160.3256 | 199 | 38 | 138 | [15 empty] | 95
| 1 | 6 | 180 | 170 | 170 | 161.7492 | 195 | 30 | 130
| Mean |  | 178.5 | 171.6667 | 176.6667 | 158.6113 | 201.1667 | 36.1667 | 143.3333
| Median |  | 180 | 170 | 180 | 158.5053 | 201 | 35.5 | 140
| Std Dev. |  | 3.7283 | 4.0825 | 5.164 | 2.2741 | 4.1191 | 4.9565 | 14.8279
| Rel.Std.Dev. |  | 2.089% | 2.378% | 2.923% | 1.434% | 2.048% | 13.7% | 10.35%
| PDM3 |  | 4.568% | 0.5652% | 3.494% | -7.083% | 17.85% | -78.81% | -16.03%
| Z-Score (Absolute) |  | 0.35 | 0 | 0.26 | 0.67 | 1.52 | 7 | 1.46
| Analytical results for Ba in OREAS 550b (Certified Value 735 ppm) | [23 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | [7 empty] | 0
0
| 1 | 1 | 738 | 734 | 709 | 744 | 684 | 775 | 740 | 710.5606 | 739.1399 | 727 | 741 | 760 | 93.5 | 694 | 676 | [7 empty] | 1
| 1 | 2 | 701 | 743 | 712 | 742 | 668 | 779 | 720 | 739.3873 | 731.9709 | 745 | 740 | 752 | 93.5 | 739 | 598 | [7 empty] | #N/A
| 1 | 3 | 714 | 734 | 736 | 760 | 689 | 773 | 750 | 724.5955 | 717.709 | 737 | 729 | 757 | 93.5 | 659 | 710 | [7 empty] | 16
| 1 | 4 | 756 | 743 | 728 | 740 | 714 | 779 | 760 | 725.8317 | 729.5947 | 746 | 735 | 751 | 93.6 | 741 | 718 | [7 empty] | 735.4886
| 1 | 5 | 737 | 752 | 750 | 767 | 623 | 779 | 770 | 712.715 | 731.6939 | 742 | 779 | 743 | 93.5 | 723 | 750 | [7 empty] | 96
| 1 | 6 | 737 | 734 | 746 | 740 | 714 | 775 | 730 | 735.7602 | 716.9664 | 728 | 759 | 795 | 93.5 | 768 | 718
| Mean |  | 730.5 | 740 | 730.1667 | 748.8333 | 682 | 776.6667 | 745 | 724.8084 | 727.8458 | 737.5 | 747.1667 | 759.6667 | 93.5167 | 720.6667 | 695
| Median |  | 737 | 738.5 | 732 | 743 | 686.5 | 777 | 745 | 725.2136 | 730.6443 | 739.5 | 740.5 | 754.5 | 93.5 | 731 | 714
| Std Dev. |  | 19.6647 | 7.3485 | 17.0929 | 11.669 | 33.9941 | 2.6583 | 18.7083 | 11.6862 | 8.7585 | 8.3606 | 18.5517 | 18.2611 | 0.040825 | 38.7539 | 53.0622
| Rel.Std.Dev. |  | 2.692% | 0.993% | 2.341% | 1.558% | 4.984% | 0.3423% | 2.511% | 1.612% | 1.203% | 1.134% | 2.483% | 2.404% | 0.04366% | 5.378% | 7.635%
| PDM3 |  | -0.6783% | 0.6134% | -0.7236% | 1.814% | -7.273% | 5.599% | 1.293% | -1.452% | -1.039% | 0.2735% | 1.588% | 3.287% | -87.29% | -2.015% | -5.505%
| Z-Score (Absolute) |  | 0 | 0.46 | 0.02 | 0.85 | 2.26 | 2.15 | 0.67 | 0.26 | 0.12 | 0.33 | 0.78 | 1.36 | 29.62 | 0.46 | 1.65
| Analytical results for Be in OREAS 550b (Certified Value 3.76 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 17 | 18 | 22 | 25 | 26 | [9 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | [9 empty] | 2
2
| 1 | 1 | 3 | 3.2 | 4 | 4 | 5 | 6 | < 5 | 3.8299 | < 5 | 4 | < 2 | 3 | 3.5 | [9 empty] | 1
| 1 | 2 | 4 | 2.4 | 4 | 3 | 4 | 6 | < 5 | 3.803 | < 5 | 4 | < 2 | 4 | 3.9 | [9 empty] | 6
| 1 | 3 | 4 | 0.9 | 4 | 4 | 4 | 6 | 5 | 3.8353 | < 5 | 4 | < 2 | 4 | 3.4 | [9 empty] | 16
| 1 | 4 | 4 | 3.7 | 4 | 4 | 5 | 6 | < 5 | 3.5368 | < 5 | 4 | < 2 | 5 | 3.5 | [9 empty] | 3.76
| 1 | 5 | 4 | 2.9 | 4 | 3 | 4 | 6 | 5 | 3.6421 | < 5 | 4 | < 2 | 4 | 3.3 | [9 empty] | 97
| 1 | 6 | 4 | 3 | 4 | 3 | 4 | 6 | < 5 | 3.6339 | < 5 | 4 | < 2 | 4 | 3.4
| Mean |  | 3.8333 | 2.6833 | 4 | 3.5 | 4.3333 | 6 | 5 | 3.7135 |  | 4 |  | 4 | 3.5
| Median |  | 4 | 2.95 | 4 | 3.5 | 4 | 6 | 5 | 3.7226 |  | 4 |  | 4 | 3.45
| Std Dev. |  | 0.40825 | 0.9704 | 0 | 0.54772 | 0.5164 | 0 | 0 | 0.12575 |  | 0 |  | 0.63246 | 0.20976
| Rel.Std.Dev. |  | 10.65% | 36.16% | 0% | 15.65% | 11.92% | 0% | 0% | 3.386% |  | 0% |  | 15.81% | 5.993%
| PDM3 |  | 1.95% | -28.64% | 6.382% | -6.915% | 15.25% | 59.57% | 32.98% | -1.237% |  | 6.382% |  | 6.382% | -6.915%
| Z-Score (Absolute) |  | 0.15 | 2.69 | 0.56 | 0.67 | 1.38 | 5.49 | NA | 0.15 | 3.14 | 0.56 | 6.84 | 0.56 | 0.67
Results from laboratory 8 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 550b (Certified Value 2.15 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 17 | 18 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 2.4 | < 2 | 2 | 1.9615 | 2.2 | 2.1 | < 5 | 2.2 | 2.1 | [13 empty] | 1
| 1 | 2 | 2.4 | 2 | 2 | 2.1159 | 2.3 | 2.1 | < 5 | 2.2 | 2.4 | [13 empty] | 7
| 1 | 3 | 2.3 | 2 | 2 | 2.0016 | 2.3 | 2.2 | < 5 | 2 | 2.2 | [13 empty] | 16
| 1 | 4 | 2.5 | 2 | 2 | 1.8533 | 2.3 | 2 | < 5 | 2.2 | 2.2 | [13 empty] | 2.1547
| 1 | 5 | 2.2 | 2 | 2.1 | 2.0433 | 2.3 | 2 | < 5 | 2.1 | 2.2 | [13 empty] | 98
| 1 | 6 | 2.2 | 2 | 2 | 1.9213 | 2.3 | 2.1 | < 5 | 2.1 | 2.4
| Mean |  | 2.3333 | 2 | 2.0167 | 1.9828 | 2.2833 | 2.0833 |  | 2.1333 | 2.25
| Median |  | 2.35 | 2 | 2 | 1.9815 | 2.3 | 2.1 |  | 2.15 | 2.2
| Std Dev. |  | 0.12111 | 0 | 0.040825 | 0.092435 | 0.040825 | 0.075277 |  | 0.08165 | 0.12247
| Rel.Std.Dev. |  | 5.19% | 0% | 2.024% | 4.662% | 1.788% | 3.613% |  | 3.827% | 5.443%
| PDM3 |  | 8.291% | -7.179% | -6.406% | -7.977% | 5.971% | -3.312% |  | -0.991% | 4.424%
| Z-Score (Absolute) |  | 0.76 | NA | 0.94 | 1.13 | 0.49 | 0.58 | 1.66 | 0.31 | 0.31
Results from laboratory 3 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 550b (Certified Value 1.35 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 1.4 | 1.387 | 1.3 | 1.3 | 1.4322 | 1.4 | 1.28 | 1.35 | 1.2906 | 1.229 | 1.258 | 1.379 | 1.379 | 1.379 | 1.3 | 1.4 | 1.3598 | 1.33 | 1.32 | 1.41 | 1.4 |  | 1
| 1 | 2 | 1.4 | 1.394 | 1.32 | 1.4 | 1.4286 | 1.4 | 1.28 | 1.31 | 1.3007 | 1.186 | 1.301 | 1.387 | 1.472 | 1.358 | 1.3 | 1.5 | 1.3579 | 1.3 | 1.33 | 1.42 | 1.4 |  | #N/A
| 1 | 3 | 1.4 | 1.73 | 1.2 | 1.3 | 1.4301 | 1.4 | 1.27 | 1.34 | 1.3093 | 1.244 | 1.251 | 1.387 | 1.458 | 1.344 | 1.3 | 1.4 | 1.4365 | 1.32 | 1.36 | 1.42 | 1.4 |  | 16
| 1 | 4 | 1.5 | 1.451 | 1.36 | 1.3 | 1.4269 | 1.5 | 1.29 | 1.31 | 1.2738 | 1.272 | 1.229 | 1.379 | 1.144 | 1.379 | 1.4 | 1.4 | 1.4201 | 1.28 | 1.33 | 1.28 | 1.4 |  | 1.3528
| 1 | 5 | 1.4 | 1.408 | 1.3 | 1.3 | 1.4314 | 1.2 | 1.29 | 1.36 | 1.289 | 1.272 | 1.244 | 1.387 | 1.372 | 1.401 | 1.4 | 1.5 | 1.337 | 1.3 | 1.35 | 1.28 | 1.4 |  | 99
| 1 | 6 | 1.4 | 1.708 | 1.34 | 1.4 | 1.4373 | 1.5 | 1.28 | 1.32 | 1.2653 | 1.229 | 1.251 | 1.387 | 1.358 | 1.379 | 1.3 | 1.5 | 1.3774 | 1.33 | 1.32 | 1.3 | 1.4
| Mean |  | 1.4167 | 1.513 | 1.3033 | 1.3333 | 1.4311 | 1.4 | 1.2817 | 1.3317 | 1.2881 | 1.2387 | 1.2557 | 1.3843 | 1.3638 | 1.3733 | 1.3333 | 1.45 | 1.3815 | 1.31 | 1.335 | 1.3517 | 1.4
| Median |  | 1.4 | 1.4295 | 1.31 | 1.3 | 1.4308 | 1.4 | 1.28 | 1.33 | 1.2898 | 1.2365 | 1.251 | 1.387 | 1.3755 | 1.379 | 1.3 | 1.45 | 1.3686 | 1.31 | 1.33 | 1.355 | 1.4
| Std Dev. |  | 0.040825 | 0.16126 | 0.055737 | 0.05164 | 0.0035944 | 0.10954 | 0.0075277 | 0.02137 | 0.016376 | 0.032284 | 0.024295 | 0.0041312 | 0.11766 | 0.019786 | 0.05164 | 0.054772 | 0.038833 | 0.02 | 0.016432 | 0.071671 | 2.4324e-16
| Rel.Std.Dev. |  | 2.882% | 10.66% | 4.277% | 3.873% | 0.2512% | 7.825% | 0.5873% | 1.605% | 1.271% | 2.606% | 1.935% | 0.2984% | 8.627% | 1.441% | 3.873% | 3.777% | 2.811% | 1.527% | 1.231% | 5.302% | 1.737e-14%
| PDM3 |  | 4.725% | 11.85% | -3.653% | -1.435% | 5.791% | 3.493% | -5.255% | -1.559% | -4.778% | -8.434% | -7.177% | 2.335% | 0.8192% | 1.522% | -1.435% | 7.189% | 2.122% | -3.16% | -1.312% | -0.08016% | 3.493%
| Z-Score (Absolute) |  | 0.91 | 2.25 | 0.67 | 0.26 | 1.11 | 0.67 | 0.98 | 0.28 | 0.89 | 1.57 | 1.34 | 0.45 | 0.17 | 0.3 | 0.26 | 1.37 | 0.42 | 0.58 | 0.23 | 0 | 0.67
| Analytical results for Cd in OREAS 550b (Indicative Value < 10 ppm) | [23 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 03 | 06 | 08 | 09 | 17 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [13 empty] | 1
1
| 1 | 1 | < 50 | < 2 | < 0.2 | < 10 | < 10 | 0.2 | < 1 | 0.6 | < 0.5 | [13 empty] | 1
| 1 | 2 | < 50 | < 2 | < 0.2 | < 10 | < 10 | < 0.2 | < 1 | 0.3 | < 0.5 | [13 empty] | 8
| 1 | 3 | < 50 | < 2 | 0.2 | < 10 | < 10 | 0.2 | < 1 | 0.7 | < 0.5 | [13 empty] | 16
| 1 | 4 | < 50 | < 2 | 0.4 | < 10 | < 10 | < 0.2 | < 1 | 0.6 | < 0.5 | [13 empty] | < 10
| 1 | 5 | < 50 | < 2 | < 0.2 | < 10 | < 10 | 0.2 | < 1 | 0.7 | < 0.5 | [13 empty] | 27
| 1 | 6 | < 50 | < 2 | 0.4 | < 10 | < 10 | 0.2 | < 1 | 0.3 | < 0.5
| Mean |  |  |  | 0.33333 |  |  | 0.2 |  | 0.53333
| Median |  |  |  | 0.4 |  |  | 0.2 |  | 0.6
| Std Dev. |  |  |  | 0.11547 |  |  | 0 |  | 0.18619
| Rel.Std.Dev. |  |  |  | 34.64% |  |  | 0% |  | 34.91%
PDM3
| Z-Score (Absolute) |  | 44.99 | 0.86 | 0.58 | 8.21 | 8.21 | 0.67 | 0.06 | 0 | 0.52
| Analytical results for Ce in OREAS 550b (Certified Value 126 ppm) | [23 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 0
0
| 1 | 1 | 116 | 118 | 128 | 123.0701 | 126 | 127 | 124.6 | 128 | 128 | [13 empty] | 1
| 1 | 2 | 105 | 121 | 125 | 125.8524 | 129 | 125 | 124.9 | 119 | 133 | [13 empty] | 9
| 1 | 3 | 101 | 116 | 125 | 123.8805 | 129 | 128 | 125.6 | 124 | 129 | [13 empty] | 16
| 1 | 4 | 127 | 120 | 141 | 123.8938 | 128 | 129 | 123.3 | 130 | 130 | [13 empty] | 126.4803
| 1 | 5 | 111 | 123 | 135 | 124.3307 | 130 | 126 | 123.4 | 129 | 130 | [13 empty] | 100
| 1 | 6 | 102 | 117 | 139 | 125.0268 | 129 | 127 | 124.2 | 132 | 126
| Mean |  | 110.3333 | 119.1667 | 132.1667 | 124.3424 | 128.5 | 127 | 124.3333 | 127 | 129.3333
| Median |  | 108 | 119 | 131.5 | 124.1122 | 129 | 127 | 124.4 | 128.5 | 129.5
| Std Dev. |  | 9.9532 | 2.6394 | 7.111 | 0.97794 | 1.3784 | 1.4142 | 0.88919 | 4.7329 | 2.3381
| Rel.Std.Dev. |  | 9.021% | 2.215% | 5.38% | 0.7865% | 1.073% | 1.114% | 0.7152% | 3.727% | 1.808%
| PDM3 |  | -12.77% | -5.782% | 4.496% | -1.69% | 1.597% | 0.4109% | -1.697% | 0.4109% | 2.256%
| Z-Score (Absolute) |  | 4.23 | 1.99 | 1.31 | 0.67 | 0.38 | 0 | 0.68 | 0 | 0.59
| Analytical results for Co in OREAS 550b (Certified Value 0.145 wt.%) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | 3
3
| 1 | 1 | 0.1484 | 0.138 | 0.141 | 0.145 | 0.14381 | 0.1397 | 0.1435 | 0.149 | 0.14914 | 0.144 | 0.139 | 0.144 | 0.141 | 0.14 | 0.14635 | 0.1472 | 0.1447 | 0.1469 | 0.146 | 0.11632 | 0.13166 | 0.1433 | 1
| 1 | 2 | 0.1527 | 0.135 | 0.147 | 0.144 | 0.1444 | 0.1333 | 0.1435 | 0.145 | 0.15358 | 0.141 | 0.149 | 0.144 | 0.142 | 0.141 | 0.14619 | 0.1523 | 0.1469 | 0.1479 | 0.144 | 0.12083 | 0.13169 | 0.1463 | 29
| 1 | 3 | 0.1503 | 0.135 | 0.148 | 0.144 | 0.14508 | 0.1295 | 0.1432 | 0.139 | 0.15488 | 0.146 | 0.142 | 0.144 | 0.147 | 0.14 | 0.14544 | 0.147 | 0.1428 | 0.1496 | 0.143 | 0.10929 | 0.1329 | 0.1376 | 16
| 1 | 4 | 0.1461 | 0.141 | 0.148 | 0.146 | 0.14401 | 0.1427 | 0.1451 | 0.146 | 0.15013 | 0.144 | 0.142 | 0.143 | 0.145 | 0.145 | 0.14825 | 0.1464 | 0.1432 | 0.15 | 0.142 | 0.11376 | 0.13194 | 0.1386 | 0.14509
| 1 | 5 | 0.1495 | 0.136 | 0.152 | 0.146 | 0.14454 | 0.1343 | 0.1442 | 0.151 | 0.148 | 0.145 | 0.142 | 0.144 | 0.141 | 0.145 | 0.14593 | 0.1434 | 0.1485 | 0.1495 | 0.142 | 0.11531 | 0.13045 | 0.1432 | 101
| 1 | 6 | 0.1495 | 0.142 | 0.152 | 0.147 | 0.14469 | 0.1405 | 0.1439 | 0.141 | 0.15064 | 0.147 | 0.143 | 0.147 | 0.146 | 0.145 | 0.14383 | 0.1457 | 0.1478 | 0.1492 | 0.144 | 0.11485 | 0.13308 | 0.1383
| Mean |  | 0.14942 | 0.13783 | 0.148 | 0.14533 | 0.14442 | 0.13667 | 0.1439 | 0.14517 | 0.15106 | 0.1445 | 0.14283 | 0.14433 | 0.14367 | 0.14267 | 0.146 | 0.147 | 0.14565 | 0.14885 | 0.1435 | 0.11506 | 0.13195 | 0.14122
| Median |  | 0.1495 | 0.137 | 0.148 | 0.1455 | 0.14447 | 0.137 | 0.1437 | 0.1455 | 0.15038 | 0.1445 | 0.142 | 0.144 | 0.1435 | 0.143 | 0.14606 | 0.1467 | 0.1458 | 0.14935 | 0.1435 | 0.11508 | 0.13181 | 0.1409
| Std Dev. |  | 0.0021729 | 0.0030605 | 0.0040497 | 0.0012111 | 0.0004627 | 0.0050714 | 0.00068411 | 0.0045789 | 0.0026476 | 0.0020736 | 0.0033116 | 0.0013663 | 0.0026583 | 0.002582 | 0.0014317 | 0.0029373 | 0.0024222 | 0.0011946 | 0.0015166 | 0.0037409 | 0.00095834 | 0.0035369
| Rel.Std.Dev. |  | 1.454% | 2.22% | 2.736% | 0.8333% | 0.3204% | 3.711% | 0.4754% | 3.154% | 1.753% | 1.435% | 2.319% | 0.9466% | 1.85% | 1.81% | 0.9806% | 1.998% | 1.663% | 0.8025% | 1.057% | 3.251% | 0.7263% | 2.505%
| PDM3 |  | 2.985% | -4.999% | 2.009% | 0.1708% | -0.4568% | -5.803% | -0.8172% | 0.05589% | 4.119% | -0.4036% | -1.552% | -0.5185% | -0.978% | -1.667% | 0.6293% | 1.32% | 0.389% | 2.595% | -1.093% | -20.7% | -9.051% | -2.667%
| Z-Score (Absolute) |  | 2.04 | 2.65 | 1.47 | 0.39 | 0.02 | 3.12 | 0.19 | 0.32 | 2.71 | 0.05 | 0.63 | 0.02 | 0.29 | 0.69 | 0.66 | 1.06 | 0.51 | 1.81 | 0.36 | 11.87 | 5.03 | 1.28
| Analytical results for Cr in OREAS 550b (Certified Value 87 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 25 | 26 |  |  |  | ppm
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  |  |  | 0
0
| 1 | 1 | 86 | 96 | 100 | 60 | 105 | 85 | 80 | 73.0171 | 70 | 70 | 140 | 140 | 140 | 85 | 84 | 89 | < 100 | 135 | 96 |  |  |  | 1
| 1 | 2 | 79 | 75 | 80 | 70 | 120 | 85 | 70 | 73.2547 | 70 | 70 | 140 | 140 | 140 | 87 | 82 | 89 | < 100 | 120 | 96 |  |  |  | #N/A
| 1 | 3 | 82 | 82 | 80 | 70 | 110 | 84 | 80 | 77.9539 | 140 | 70 | 140 | 140 | 140 | 85 | 84 | 93 | < 100 | 141 | 90 |  |  |  | 16
| 1 | 4 | 92 | 96 | 90 | 70 | 110 | 85 | 80 | 70.7819 | 140 | 70 | 140 | 480 | 140 | 88 | 85 | 89 | < 100 | 149 | 96 |  |  |  | 87.1448
| 1 | 5 | 80 | 75 | 90 | 70 | 95 | 84 | 70 | 71.9426 | 140 | 140 | 140 | 210 | 140 | 85 | 90 | 90 | < 100 | 135 | 94 |  |  |  | 102
| 1 | 6 | 83 | 82 | 80 | 70 | 115 | 86 | 80 | 77.6994 | 70 | 140 | 140 | 210 | 140 | 86 | 98 | 90 | < 100 | 124 | 99
| Mean |  | 83.6667 | 84.3333 | 86.6667 | 68.3333 | 109.1667 | 84.8333 | 76.6667 | 74.1083 | 105 | 93.3333 | 140 | 220 | 140 | 86 | 87.1667 | 90 |  | 134 | 95.1667
| Median |  | 82.5 | 82 | 85 | 70 | 110 | 85 | 80 | 73.1359 | 105 | 70 | 140 | 175 | 140 | 85.5 | 84.5 | 89.5 |  | 135 | 96
| Std Dev. |  | 4.761 | 9.5638 | 8.165 | 4.0825 | 8.612 | 0.75277 | 5.164 | 3.0119 | 38.3406 | 36.1478 | 0 | 131.9091 | 0 | 1.2649 | 5.947 | 1.5492 |  | 10.6958 | 2.9944
| Rel.Std.Dev. |  | 5.69% | 11.34% | 9.421% | 5.974% | 7.889% | 0.8874% | 6.736% | 4.064% | 36.51% | 38.73% | 0% | 59.96% | 0% | 1.471% | 6.823% | 1.721% |  | 7.982% | 3.147%
| PDM3 |  | -3.991% | -3.226% | -0.5487% | -21.59% | 25.27% | -2.652% | -12.02% | -14.96% | 20.49% | 7.101% | 60.65% | 152.5% | 60.65% | -1.314% | 0.02504% | 3.276% |  | 53.77% | 9.205%
| Z-Score (Absolute) |  | 0.22 | 0.18 | 0.03 | 1.21 | 1.41 | 0.15 | 0.67 | 0.84 | 0.99 | 0.26 | 3.19 | 8.32 | 3.19 | 0.07 | 0 | 0.18 | 2.39 | 3.01 | 0.51
| Analytical results for Cs in OREAS 550b (Certified Value 2.26 ppm) | [23 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 14 | 15 | 17 | 18 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 2
3
| 1 | 1 | 2.3 | 2.5 | 2.33 | 2.2 | 1.8601 | 2.4 | < 5 | 2.472 | 2.2 | 2.1 | 2.6 | 2.1 | [10 empty] | 1
| 1 | 2 | 2.2 | 2.4 | 2.29 | 2.2 | 1.9894 | 2.4 | < 5 | 2.3216 | 2.2 | 2.1 | 2.3 | 2.3 | [10 empty] | 12
| 1 | 3 | 2.3 | 2.2 | 2.2 | 2.2 | 1.9571 | 2.2 | < 5 | 2.4239 | 2.2 | 2 | 2.2 | 2.2 | [10 empty] | 16
| 1 | 4 | 2.6 | 2.5 | 2.17 | 2.4 | 1.9491 | 2.1 | < 5 | 2.2762 | 2.2 | 2.1 | 2.2 | 2.2 | [10 empty] | 2.2573
| 1 | 5 | 2.3 | 2.6 | 2.42 | 2.3 | 1.8634 | 2.5 | < 5 | 2.3814 | 2.1 | 2.2 | 2.2 | 2.3 | [10 empty] | 103
| 1 | 6 | 2.2 | 2 | 2.17 | 2.3 | 1.9568 | 2.4 | < 5 | 2.3399 | 2.2 | 2.2 | 2 | 2.3
| Mean |  | 2.3167 | 2.3667 | 2.2633 | 2.2667 | 1.9293 | 2.3333 |  | 2.3692 | 2.1833 | 2.1167 | 2.25 | 2.2333
| Median |  | 2.3 | 2.45 | 2.245 | 2.25 | 1.9529 | 2.4 |  | 2.3607 | 2.2 | 2.1 | 2.2 | 2.25
| Std Dev. |  | 0.1472 | 0.22509 | 0.10113 | 0.08165 | 0.054152 | 0.15055 |  | 0.071405 | 0.040825 | 0.075277 | 0.19748 | 0.08165
| Rel.Std.Dev. |  | 6.354% | 9.511% | 4.468% | 3.602% | 2.807% | 6.452% |  | 3.014% | 1.87% | 3.556% | 8.777% | 3.656%
| PDM3 |  | 2.632% | 4.847% | 0.2695% | 0.4172% | -14.53% | 3.371% |  | 4.958% | -3.275% | -6.228% | -0.3212% | -1.06%
| Z-Score (Absolute) |  | 0.46 | 0.91 | 0.01 | 0.01 | 3.02 | 0.61 | 2.11 | 0.94 | 0.73 | 1.33 | 0.13 | 0.28
| Analytical results for Cu in OREAS 550b (Certified Value 1.83 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 1.8289 | 1.858 | 1.84 | 1.81 | 1.7526 | 1.8952 | 1.87 | 1.826 | 1.8245 | 1.785 | 1.77 | 1.825 | 1.83 | 1.83 | 1.88 | 1.8517 | 1.8009 | 1.802 | 1.9153 | 1.6147 | 2.0208 |  | 1
| 1 | 2 | 1.8212 | 1.844 | 1.85 | 1.84 | 1.7592 | 1.847 | 1.86 | 1.768 | 1.8493 | 1.78 | 1.83 | 1.905 | 1.815 | 1.825 | 1.83 | 1.8546 | 1.7878 | 1.773 | 1.9244 | 1.6117 | 1.9753 |  | 30
| 1 | 3 | 1.8281 | 1.843 | 1.79 | 1.8 | 1.7603 | 1.9141 | 1.86 | 1.809 | 1.8352 | 1.78 | 1.815 | 1.85 | 1.825 | 1.81 | 1.85 | 1.8003 | 1.816 | 1.773 | 1.9644 | 1.6331 | 1.9575 |  | 16
| 1 | 4 | 1.8112 | 1.889 | 1.9 | 1.8 | 1.7556 | 2.0489 | 1.86 | 1.793 | 1.8587 | 1.78 | 1.77 | 1.795 | 1.77 | 1.85 | 1.83 | 1.823 | 1.8722 | 1.745 | 1.9426 | 1.6382 | 1.9652 |  | 1.8305
| 1 | 5 | 1.8409 | 1.877 | 1.86 | 1.83 | 1.7579 | 1.8446 | 1.86 | 1.864 | 1.8587 | 1.8 | 1.795 | 1.77 | 1.8 | 1.87 | 1.83 | 1.9144 | 1.7717 | 1.755 | 1.9345 | 1.6275 | 1.9753 |  | 104
| 1 | 6 | 1.8285 | 1.908 | 1.91 | 1.83 | 1.7524 | 2.0125 | 1.85 | 1.756 | 1.8256 | 1.825 | 1.785 | 1.815 | 1.79 | 1.845 | 1.82 | 1.88 | 1.8131 | 1.78 | 1.9496 | 1.6172 | 2.0716
| Mean |  | 1.8265 | 1.8698 | 1.8583 | 1.8183 | 1.7563 | 1.9271 | 1.86 | 1.8027 | 1.842 | 1.7917 | 1.7942 | 1.8267 | 1.805 | 1.8383 | 1.84 | 1.854 | 1.8103 | 1.7713 | 1.9385 | 1.6237 | 1.9943
| Median |  | 1.8283 | 1.8675 | 1.855 | 1.82 | 1.7568 | 1.9046 | 1.86 | 1.801 | 1.8423 | 1.7825 | 1.79 | 1.82 | 1.8075 | 1.8375 | 1.83 | 1.8531 | 1.807 | 1.773 | 1.9385 | 1.6223 | 1.9753
| Std Dev. |  | 0.0098169 | 0.026088 | 0.043551 | 0.017224 | 0.0033673 | 0.085484 | 0.0063246 | 0.039546 | 0.015679 | 0.018074 | 0.024376 | 0.046975 | 0.022804 | 0.021134 | 0.021909 | 0.040434 | 0.034515 | 0.019906 | 0.017702 | 0.010791 | 0.043819
| Rel.Std.Dev. |  | 0.5375% | 1.395% | 2.344% | 0.9472% | 0.1917% | 4.436% | 0.34% | 2.194% | 0.8512% | 1.009% | 1.359% | 2.572% | 1.263% | 1.15% | 1.191% | 2.181% | 1.907% | 1.124% | 0.9132% | 0.6646% | 2.197%
| PDM3 |  | -0.2192% | 2.15% | 1.522% | -0.6635% | -4.05% | 5.276% | 1.613% | -1.519% | 0.6297% | -2.12% | -1.984% | -0.2083% | -1.392% | 0.4291% | 0.5201% | 1.285% | -1.103% | -3.231% | 5.899% | -11.29% | 8.949%
| Z-Score (Absolute) |  | 0 | 0.92 | 0.67 | 0.18 | 1.5 | 2.14 | 0.71 | 0.51 | 0.33 | 0.75 | 0.69 | 0 | 0.46 | 0.25 | 0.28 | 0.58 | 0.35 | 1.18 | 2.38 | 4.32 | 3.57
| Analytical results for Dy in OREAS 550b (Certified Value 5.78 ppm) | [23 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 5.6 | 6.11 | 5.455 | 5.93 | 6 | 4.9 | 6.14 | 5.58 | [14 empty] | 1
| 1 | 2 | 5.6 | 6.02 | 5.4457 | 5.82 | 6 | 4.7 | 5.32 | 5.67 | [14 empty] | 13
| 1 | 3 | 5 | 6.17 | 5.4841 | 5.75 | 6 | 4.5 | 5.84 | 5.53 | [14 empty] | 16
| 1 | 4 | 5.9 | 7.06 | 5.4783 | 5.6 | 6 | 4.7 | 5.88 | 5.68 | [14 empty] | 5.7834
| 1 | 5 | 5.3 | 6.1 | 5.3911 | 5.86 | 6 | 4.7 | 6.42 | 5.68 | [14 empty] | 105
| 1 | 6 | 5.6 | 6.68 | 5.4446 | 5.73 | 6 | 4.8 | 5.96 | 5.52
| Mean |  | 5.5 | 6.3567 | 5.4498 | 5.7817 | 6 | 4.7167 | 5.9267 | 5.61
| Median |  | 5.6 | 6.14 | 5.4504 | 5.785 | 6 | 4.7 | 5.92 | 5.625
| Std Dev. |  | 0.30984 | 0.41812 | 0.033196 | 0.11514 | 0 | 0.13292 | 0.36544 | 0.075895
| Rel.Std.Dev. |  | 5.633% | 6.578% | 0.6091% | 1.991% | 0% | 2.818% | 6.166% | 1.353%
| PDM3 |  | -4.901% | 9.911% | -5.769% | -0.03077% | 3.744% | -18.45% | 2.476% | -2.999%
| Z-Score (Absolute) |  | 0.55 | 1.87 | 0.7 | 0.24 | 0.86 | 2.77 | 0.65 | 0.24
| Analytical results for Er in OREAS 550b (Certified Value 3.14 ppm) | [23 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
2
| 1 | 1 | 3.2 | 3.43 | 3.2564 | 3.38 | 3 | 2.8 | 3.58 | 2.8 | [14 empty] | 1
| 1 | 2 | 3.2 | 3.38 | 3.2838 | 3.28 | 3 | 2.7 | 3.2 | 3.03 | [14 empty] | 14
| 1 | 3 | 2.4 | 3.48 | 3.2875 | 3.38 | 3 | 2.6 | 3.19 | 2.91 | [14 empty] | 16
| 1 | 4 | 3.5 | 3.75 | 3.1783 | 3.35 | 3 | 2.7 | 3.57 | 2.77 | [14 empty] | 3.1387
| 1 | 5 | 2.9 | 3.49 | 3.3256 | 3.53 | 3 | 2.9 | 3.31 | 2.89 | [14 empty] | 106
| 1 | 6 | 2.3 | 3.88 | 3.2239 | 3.39 | 3 | 2.7 | 3.41 | 2.82
| Mean |  | 2.9167 | 3.5683 | 3.2592 | 3.385 | 3 | 2.7333 | 3.3767 | 2.87
| Median |  | 3.05 | 3.485 | 3.2701 | 3.38 | 3 | 2.7 | 3.36 | 2.855
| Std Dev. |  | 0.47924 | 0.19934 | 0.052177 | 0.08167 | 0 | 0.10328 | 0.1734 | 0.094868
| Rel.Std.Dev. |  | 16.43% | 5.586% | 1.601% | 2.413% | 0% | 3.779% | 5.135% | 3.306%
| PDM3 |  | -7.073% | 13.69% | 3.842% | 7.849% | -4.418% | -12.91% | 7.583% | -8.56%
| Z-Score (Absolute) |  | 0.57 | 1.18 | 0.35 | 0.69 | 0.35 | 1.06 | 0.66 | 0.7
| Analytical results for Eu in OREAS 550b (Certified Value 0.92 ppm) | [23 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 0.9 | 1.07 | 0.93018 | 1.01 | < 2 | 0.9 | 0.88 | 0.94 | [14 empty] | 1
| 1 | 2 | 0.9 | 0.93 | 0.88452 | 1 | < 2 | 0.9 | 0.82 | 1.04 | [14 empty] | 15
| 1 | 3 | 0.8 | 0.93 | 0.86691 | 1.04 | < 2 | 0.9 | 0.92 | 1.01 | [14 empty] | 16
| 1 | 4 | 0.9 | 1.06 | 0.93311 | 1.03 | < 2 | 0.9 | 1.02 | 1 | [14 empty] | 0.92213
| 1 | 5 | 0.9 | 1.04 | 0.89074 | 1.01 | < 2 | 0.9 | 0.85 | 1.02 | [14 empty] | 107
| 1 | 6 | 1 | 1.15 | 0.96829 | 0.98 | < 2 | 0.9 | 0.99 | 0.9
| Mean |  | 0.9 | 1.03 | 0.91229 | 1.0117 |  | 0.9 | 0.91333 | 0.985
| Median |  | 0.9 | 1.05 | 0.91046 | 1.01 |  | 0.9 | 0.9 | 1.005
| Std Dev. |  | 0.063246 | 0.086023 | 0.037906 | 0.02137 |  | 0 | 0.078909 | 0.053572
| Rel.Std.Dev. |  | 7.027% | 8.352% | 4.155% | 2.112% |  | 0% | 8.64% | 5.439%
| PDM3 |  | -2.399% | 11.7% | -1.066% | 9.71% |  | -2.399% | -0.9534% | 6.818%
| Z-Score (Absolute) |  | 0.67 | 1.11 | 0.51 | 0.86 | 0.7 | 0.67 | 0.49 | 0.49
| Analytical results for Fe in OREAS 550b (Certified Value 2.74 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 2.77 | 2.66 | 2.81 | 2.76 | 2.7851 | 2.64 | 2.72 | 2.66 | 2.7157 | 2.658 | 2.49 | 2.742 | 2.812 | 2.721 | 2.9055 | 2.7 | 2.75 | 2.8456 | 2.63 | 2.43 | 2.61 | 2.88 | 1
| 1 | 2 | 2.82 | 2.6 | 2.78 | 2.76 | 2.7828 | 2.61 | 2.72 | 2.72 | 2.8076 | 2.658 | 2.679 | 2.791 | 2.784 | 2.679 | 2.9465 | 2.76 | 2.76 | 2.8204 | 2.61 | 2.44 | 2.46 | 2.84 | #N/A
| 1 | 3 | 2.77 | 2.64 | 2.78 | 2.75 | 2.7898 | 2.64 | 2.72 | 2.65 | 2.8016 | 2.742 | 2.588 | 2.77 | 2.798 | 2.637 | 2.899 | 2.73 | 2.7 | 2.8944 | 2.61 | 2.46 | 2.69 | 2.68 | 16
| 1 | 4 | 2.71 | 2.69 | 2.79 | 2.78 | 2.7809 | 2.75 | 2.72 | 2.71 | 2.8007 | 2.728 | 2.602 | 2.77 | 2.819 | 2.693 | 2.8973 | 2.77 | 2.71 | 2.9405 | 2.56 | 2.45 | 2.82 | 2.81 | 2.7424
| 1 | 5 | 2.77 | 2.66 | 2.79 | 2.79 | 2.7876 | 2.54 | 2.71 | 2.64 | 2.7199 | 2.721 | 2.602 | 2.777 | 2.714 | 2.763 | 2.8532 | 2.75 | 2.88 | 2.8588 | 2.58 | 2.43 | 2.88 | 2.83 | 108
| 1 | 6 | 2.75 | 2.74 | 2.79 | 2.78 | 2.78 | 2.78 | 2.71 | 2.67 | 2.7559 | 2.77 | 2.588 | 2.833 | 2.777 | 2.728 | 2.8243 | 2.69 | 2.81 | 2.8718 | 2.63 | 2.44 | 2.61 | 2.99
| Mean |  | 2.765 | 2.665 | 2.79 | 2.77 | 2.7844 | 2.66 | 2.7167 | 2.675 | 2.7669 | 2.7128 | 2.5915 | 2.7805 | 2.784 | 2.7035 | 2.8876 | 2.7333 | 2.7683 | 2.8719 | 2.6033 | 2.4417 | 2.6783 | 2.8383
| Median |  | 2.77 | 2.66 | 2.79 | 2.77 | 2.784 | 2.64 | 2.72 | 2.665 | 2.7783 | 2.7245 | 2.595 | 2.7735 | 2.791 | 2.707 | 2.8981 | 2.74 | 2.755 | 2.8653 | 2.61 | 2.44 | 2.65 | 2.835
| Std Dev. |  | 0.035637 | 0.047223 | 0.010954 | 0.015492 | 0.0038281 | 0.089666 | 0.005164 | 0.032711 | 0.042311 | 0.04567 | 0.060338 | 0.03027 | 0.037826 | 0.043771 | 0.042904 | 0.03266 | 0.067355 | 0.04178 | 0.028048 | 0.01169 | 0.15355 | 0.10068
| Rel.Std.Dev. |  | 1.289% | 1.772% | 0.3926% | 0.5593% | 0.1375% | 3.371% | 0.1901% | 1.223% | 1.529% | 1.683% | 2.328% | 1.089% | 1.359% | 1.619% | 1.486% | 1.195% | 2.433% | 1.455% | 1.077% | 0.4788% | 5.733% | 3.547%
| PDM3 |  | 0.8258% | -2.821% | 1.737% | 1.008% | 1.533% | -3.003% | -0.9367% | -2.456% | 0.895% | -1.076% | -5.501% | 1.391% | 1.519% | -1.417% | 5.297% | -0.3289% | 0.9474% | 4.725% | -5.069% | -10.96% | -2.334% | 3.5%
| Z-Score (Absolute) |  | 0.25 | 1.31 | 0.64 | 0.32 | 0.55 | 1.39 | 0.51 | 1.15 | 0.28 | 0.57 | 2.46 | 0.48 | 0.54 | 0.71 | 2.16 | 0.25 | 0.3 | 1.91 | 2.27 | 4.79 | 1.1 | 1.39
| Analytical results for Ga in OREAS 550b (Certified Value 23 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 1
2
| 1 | 1 | 23 | 23.1 | 23 | 23 | 17 | 21.8593 | 22.7445 | 23 | 21 | 24.4 | 22 | 21 | [10 empty] | 1
| 1 | 2 | 23 | 22.3 | 23 | 22 | 18 | 25.0769 | 24.1313 | 23 | 23 | 24.5 | 24 | 22 | [10 empty] | 17
| 1 | 3 | 22 | 25.6 | 22 | 22 | 18 | 25.0566 | 22.946 | 22 | 21 | 25 | 21 | 21 | [10 empty] | 16
| 1 | 4 | 24 | 26.9 | 23 | 23 | 17 | 23.7331 | 23.3871 | 23 | 23 | 24.2 | 24 | 21 | [10 empty] | 23.0023
| 1 | 5 | 23 | 24.3 | 23 | 22 | 18 | 21.9013 | 23.4154 | 22 | 22 | 23.9 | 24 | 21 | [10 empty] | 109
| 1 | 6 | 22 | 25.5 | 23 | 23 | 18 | 23.4394 | 22.9354 | 22 | 24 | 24.5 | 22 | 21
| Mean |  | 22.8333 | 24.6167 | 22.8333 | 22.5 | 17.6667 | 23.5111 | 23.2599 | 22.5 | 22.3333 | 24.4167 | 22.8333 | 21.1667
| Median |  | 23 | 24.9 | 23 | 22.5 | 18 | 23.5863 | 23.1665 | 22.5 | 22.5 | 24.45 | 23 | 21
| Std Dev. |  | 0.75277 | 1.7163 | 0.40825 | 0.54772 | 0.5164 | 1.4293 | 0.5038 | 0.54772 | 1.2111 | 0.3656 | 1.3292 | 0.40825
| Rel.Std.Dev. |  | 3.297% | 6.972% | 1.788% | 2.434% | 2.923% | 6.079% | 2.166% | 2.434% | 5.423% | 1.497% | 5.821% | 1.929%
| PDM3 |  | -0.7347% | 7.018% | -0.7347% | -2.184% | -23.2% | 2.212% | 1.12% | -2.184% | -2.908% | 6.149% | -0.7347% | -7.98%
| Z-Score (Absolute) |  | 0 | 2.6 | 0 | 0.49 | 7.52 | 0.99 | 0.62 | 0.49 | 0.73 | 2.3 | 0 | 2.43
| Analytical results for Gd in OREAS 550b (Certified Value 6.12 ppm) | [23 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 22 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [15 empty] | 2
3
| 1 | 1 | 6.5 | 6.41 | 5.7029 | 6.06 | 5.7 | 6.35 | 5.96 | [15 empty] | 1
| 1 | 2 | 6.4 | 6.48 | 6.0726 | 5.85 | 5.5 | 5.85 | 6.01 | [15 empty] | 18
| 1 | 3 | 4.8 | 6.33 | 6.2949 | 6.03 | 5.7 | 5.85 | 6.21 | [15 empty] | 16
| 1 | 4 | 6.4 | 6.89 | 6.2002 | 5.72 | 5.6 | 6.04 | 6.1 | [15 empty] | 6.1183
| 1 | 5 | 6.3 | 7.06 | 6.2595 | 6.01 | 5.7 | 6.2 | 6.19 | [15 empty] | 110
| 1 | 6 | 6.8 | 7.35 | 6.0355 | 5.95 | 5.7 | 5.7 | 5.74
| Mean |  | 6.2 | 6.7533 | 6.0943 | 5.9367 | 5.65 | 5.9983 | 6.035
| Median |  | 6.4 | 6.685 | 6.1364 | 5.98 | 5.7 | 5.945 | 6.055
| Std Dev. |  | 0.70711 | 0.41001 | 0.21717 | 0.12956 | 0.083666 | 0.24457 | 0.17444
| Rel.Std.Dev. |  | 11.4% | 6.071% | 3.564% | 2.182% | 1.481% | 4.077% | 2.891%
| PDM3 |  | 1.335% | 10.38% | -0.3933% | -2.969% | -7.654% | -1.961% | -1.362%
| Z-Score (Absolute) |  | 1.13 | 4.93 | 0.41 | 0.67 | 2.64 | 0.25 | 0
| Analytical results for Ge in OREAS 550b (Certified Value 2.01 ppm) | [23 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 17 | 18 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 2 | 1.7 | 2 | 2 | 1.3823 | 2 | 2 | < 5 | 3 | [13 empty] | 1
| 1 | 2 | 2 | 2.4 | 2 | 2 | 1.6972 | 2 | 2 | < 5 | 4 | [13 empty] | 2
| 1 | 3 | 1 | 1.7 | 2 | 2 | 1.4433 | 2 | 2 | < 5 | 3 | [13 empty] | 16
| 1 | 4 | 2 | 2.3 | 2 | 2 | 1.6834 | 2 | 2 | < 5 | 4 | [13 empty] | 2.01
| 1 | 5 | 2 | 2.2 | 2 | 2 | 1.4583 | 2 | 2 | < 5 | 3 | [13 empty] | 111
| 1 | 6 | 2 | 2.8 | 2 | 2 | 1.6407 | 2 | 2 | < 5 | 3
| Mean |  | 1.8333 | 2.1833 | 2 | 2 | 1.5509 | 2 | 2 |  | 3.3333
| Median |  | 2 | 2.25 | 2 | 2 | 1.5495 | 2 | 2 |  | 3
| Std Dev. |  | 0.40825 | 0.42622 | 0 | 0 | 0.13828 | 0 | 0 |  | 0.5164
| Rel.Std.Dev. |  | 22.27% | 19.52% | 0% | 0% | 8.916% | 0% | 0% |  | 15.49%
| PDM3 |  | -8.789% | 8.624% | -0.4975% | -0.4975% | -22.84% | -0.4975% | -0.4975% |  | 65.84%
| Z-Score (Absolute) |  | 0.67 | 0.74 | 0 | 0 | 1.82 | 0 | 0 | 2.02 | 5.39
| Analytical results for Hf in OREAS 550b (Indicative Value 6.29 ppm) | [23 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 03 | 15 | 17 | 22 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | [18 empty] | 2
2
| 1 | 1 | < 10 | 4.4743 | 5 | 7.2 | [18 empty] | 1
| 1 | 2 | 10 | 4.6912 | 5 | 7.1 | [18 empty] | 22
| 1 | 3 | 10 | 4.7663 | 5 | 7.1 | [18 empty] | 16
| 1 | 4 | 10 | 4.6459 | 5 | 7.3 | [18 empty] | 6.291
| 1 | 5 | < 10 | 4.5964 | 5 | 7.4 | [18 empty] | 28
| 1 | 6 | 10 | 4.5105 | 5 | 7.2
| Mean |  | 10 | 4.6141 | 5 | 7.2167
| Median |  | 10 | 4.6212 | 5 | 7.2
| Std Dev. |  | 0 | 0.1102 | 0 | 0.1169
| Rel.Std.Dev. |  | 0% | 2.388% | 0% | 1.62%
| PDM3 |  | 58.96% | -26.66% | -20.52% | 14.71%
| Z-Score (Absolute) |  | 1.15 | 0.77 | 0.57 | 0.57
| Analytical results for Hg in OREAS 550b (Indicative Value < 5 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | Lab | [21 empty] | 1
| No | No | 22 | [21 empty] | ppm
|  |  | PF*OES | [21 empty] | 2
2
| 1 | 1 | < 5 | [21 empty] | 1
| 1 | 2 | < 5 | [21 empty] | 1
| 1 | 3 | < 5 | [21 empty] | 16
| 1 | 4 | < 5 | [21 empty] | < 5
| 1 | 5 | < 5 | [21 empty] | 26
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 550b (Certified Value 1.14 ppm) | [23 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 22 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [15 empty] | 2
3
| 1 | 1 | 1.2 | 1.15 | 1.1452 | 1.24 | 0.9 | 1.21 | 1.09 | [15 empty] | 1
| 1 | 2 | 1.1 | 1.16 | 1.1124 | 1.2 | 0.9 | 1.1 | 1.11 | [15 empty] | 5
| 1 | 3 | 1 | 1.14 | 1.0956 | 1.18 | 0.9 | 1.09 | 1.09 | [15 empty] | 16
| 1 | 4 | 1.1 | 1.29 | 1.1266 | 1.18 | 0.9 | 1.18 | 1.14 | [15 empty] | 1.1428
| 1 | 5 | 1 | 1.27 | 1.0687 | 1.2 | 0.9 | 1.2 | 1.11 | [15 empty] | 112
| 1 | 6 | 1.1 | 1.29 | 1.1012 | 1.19 | 0.9 | 1.16 | 1.02
| Mean |  | 1.0833 | 1.2167 | 1.1083 | 1.1983 | 0.9 | 1.1567 | 1.0933
| Median |  | 1.1 | 1.215 | 1.1068 | 1.195 | 0.9 | 1.17 | 1.1
| Std Dev. |  | 0.075277 | 0.073666 | 0.026442 | 0.022286 | 0 | 0.050859 | 0.040332
| Rel.Std.Dev. |  | 6.949% | 6.055% | 2.386% | 1.86% | 0% | 4.397% | 3.689%
| PDM3 |  | -5.201% | 6.467% | -3.018% | 4.862% | -21.24% | 1.216% | -4.326%
| Z-Score (Absolute) |  | 0.35 | 1.51 | 0 | 1.25 | 2.9 | 0.67 | 0.21
| Analytical results for In in OREAS 550b (Certified Value 0.98 ppm) | [23 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 17 | 18 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
2
| 1 | 1 | 1 | 0.9 | 1 | 1.1 | 0.78496 | 0.9 | 1 | 0.98 | 0.9 | [13 empty] | 1
| 1 | 2 | 0.9 | 1.2 | 1 | 1.1 | 0.78405 | 1 | 0.9 | 0.96 | 1 | [13 empty] | 6
| 1 | 3 | 0.9 | 1.2 | 1.1 | 1 | 0.92622 | 1 | 0.9 | 0.95 | 1 | [13 empty] | 16
| 1 | 4 | 0.9 | 1 | 1.1 | 1.1 | 0.79254 | 0.9 | 1.1 | 0.99 | 1 | [13 empty] | 0.98272
| 1 | 5 | 0.9 | 1.2 | 1 | 1 | 0.85756 | 0.9 | 1.1 | 1.02 | 1.1 | [13 empty] | 113
| 1 | 6 | 1.1 | 1 | 1 | 1.1 | 0.89172 | 0.9 | 0.9 | 0.83 | 1
| Mean |  | 0.95 | 1.0833 | 1.0333 | 1.0667 | 0.83951 | 0.93333 | 0.98333 | 0.955 | 1
| Median |  | 0.9 | 1.1 | 1 | 1.1 | 0.82505 | 0.9 | 0.95 | 0.97 | 1
| Std Dev. |  | 0.083666 | 0.13292 | 0.05164 | 0.05164 | 0.061362 | 0.05164 | 0.098319 | 0.065955 | 0.063246
| Rel.Std.Dev. |  | 8.807% | 12.27% | 4.997% | 4.841% | 7.309% | 5.533% | 9.999% | 6.906% | 6.325%
| PDM3 |  | -3.33% | 10.24% | 5.15% | 8.542% | -14.57% | -5.026% | 0.0621% | -2.821% | 1.758%
| Z-Score (Absolute) |  | 0.45 | 1.35 | 0.67 | 1.12 | 1.94 | 0.67 | 0 | 0.38 | 0.22
| Analytical results for K in OREAS 550b (Certified Value 2.03 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 2.02 | 2.05 | 2.1 | 2.23 | 2.2 | 2.11 | 1.93 | 1.9195 | 2.034 | 1.959 | 2.042 | 2.017 | 2.084 | 1.9 | 2.03 | 1.923 | 2.04 | 0.21 | 1.85 | 2.1 |  |  | 1
| 1 | 2 | 2.06 | 2.001 | 2 | 2.31 | 2.2 | 2.1 | 2.02 | 1.9733 | 2.001 | 2.017 | 2.059 | 2.092 | 2.092 | 2 | 2.04 | 1.9631 | 2.03 | 0.22 | 1.74 | 2.1 |  |  | #N/A
| 1 | 3 | 2.01 | 1.976 | 2 | 2.27 | 2.2 | 2.09 | 1.96 | 2.0214 | 2.034 | 1.959 | 2.1 | 2.133 | 2.084 | 2 | 1.99 | 2.0274 | 2.02 | 0.21 | 1.93 | 2.1 |  |  | 16
| 1 | 4 | 1.97 | 2.059 | 2 | 2.25 | 2.5 | 2.09 | 1.94 | 1.9268 | 2.059 | 1.992 | 2.042 | 2.034 | 2.067 | 2 | 2.02 | 2.0232 | 1.98 | 0.22 | 2.18 | 2.2 |  |  | 2.0317
| 1 | 5 | 2.02 | 2.009 | 2 | 2.35 | 2 | 2.1 | 1.98 | 1.952 | 2.05 | 2.009 | 2.034 | 2.067 | 2.092 | 2 | 2.13 | 1.9938 | 2 | 0.21 | 2.19 | 2.1 |  |  | 114
| 1 | 6 | 2 | 1.992 | 2 | 2.29 | 2.4 | 2.1 | 2.02 | 2.0011 | 2.059 | 2.001 | 2.075 | 2.067 | 2.092 | 2 | 2.09 | 1.9563 | 2.06 | 0.22 | 2.07 | 2.2
| Mean |  | 2.0133 | 2.0145 | 2.0167 | 2.2833 | 2.25 | 2.0983 | 1.975 | 1.9657 | 2.0395 | 1.9895 | 2.0587 | 2.0683 | 2.0852 | 1.9833 | 2.05 | 1.9811 | 2.0217 | 0.215 | 1.9933 | 2.1333
| Median |  | 2.015 | 2.005 | 2 | 2.28 | 2.2 | 2.1 | 1.97 | 1.9626 | 2.042 | 1.9965 | 2.0505 | 2.067 | 2.088 | 2 | 2.035 | 1.9785 | 2.025 | 0.215 | 2 | 2.1
| Std Dev. |  | 0.029439 | 0.032989 | 0.040825 | 0.043205 | 0.17607 | 0.0075277 | 0.038859 | 0.040603 | 0.02197 | 0.025042 | 0.025057 | 0.041375 | 0.0097245 | 0.040825 | 0.05099 | 0.040968 | 0.028577 | 0.0054772 | 0.18338 | 0.05164
| Rel.Std.Dev. |  | 1.462% | 1.638% | 2.024% | 1.892% | 7.825% | 0.3587% | 1.968% | 2.066% | 1.077% | 1.259% | 1.217% | 2% | 0.4664% | 2.058% | 2.487% | 2.068% | 1.414% | 2.548% | 9.199% | 2.421%
| PDM3 |  | -0.9022% | -0.8448% | -0.7381% | 12.39% | 10.75% | 3.282% | -2.789% | -3.248% | 0.3857% | -2.075% | 1.329% | 1.805% | 2.633% | -2.379% | 0.9026% | -2.487% | -0.492% | -89.42% | -1.887% | 5.004%
| Z-Score (Absolute) |  | 0.1 | 0.08 | 0.04 | 4.59 | 4.01 | 1.38 | 0.77 | 0.93 | 0.35 | 0.51 | 0.69 | 0.86 | 1.14 | 0.62 | 0.54 | 0.66 | 0.04 | 31.34 | 0.45 | 1.98
| Analytical results for La in OREAS 550b (Certified Value 66 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 15 | 17 | 19 | 22 | 25 | 26 | [11 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | [11 empty] | 0
1
| 1 | 1 | 63 | 64.3 | 66.6 | 65 | 70 | 63.0491 | 66.5 | 66 | 65.7 | 68.1 | 73 | [11 empty] | 1
| 1 | 2 | 64 | 67.1 | 65.6 | 65 | 60 | 62.9875 | 68.7 | 66 | 65.4 | 62.6 | 72 | [11 empty] | 7
| 1 | 3 | 64 | 64.5 | 65.6 | 65 | 60 | 59.834 | 68.7 | 67 | 65.8 | 65.1 | 71 | [11 empty] | 16
| 1 | 4 | 68 | 68.1 | 74.2 | 65 | 60 | 60.6972 | 68.6 | 66 | 64.7 | 69.1 | 71 | [11 empty] | 65.9656
| 1 | 5 | 67 | 68.2 | 71.1 | 65 | 70 | 62.5663 | 68.8 | 66 | 64.8 | 68.2 | 72 | [11 empty] | 115
| 1 | 6 | 66 | 70.5 | 73.2 | 65 | 60 | 60.8243 | 69.3 | 67 | 65.3 | 68.4 | 73
| Mean |  | 65.3333 | 67.1167 | 69.3833 | 65 | 63.3333 | 61.6597 | 68.4333 | 66.3333 | 65.2833 | 66.9167 | 72
| Median |  | 65 | 67.6 | 68.85 | 65 | 60 | 61.6953 | 68.7 | 66 | 65.35 | 68.15 | 72
| Std Dev. |  | 1.9664 | 2.3819 | 3.9265 | 0 | 5.164 | 1.3764 | 0.97912 | 0.5164 | 0.4535 | 2.5278 | 0.89443
| Rel.Std.Dev. |  | 3.01% | 3.549% | 5.659% | 0% | 8.154% | 2.232% | 1.431% | 0.7785% | 0.6947% | 3.778% | 1.242%
| PDM3 |  | -0.9585% | 1.745% | 5.181% | -1.464% | -3.99% | -6.527% | 3.741% | 0.5574% | -1.034% | 1.442% | 9.148%
| Z-Score (Absolute) |  | 0.51 | 0.4 | 1.54 | 0.67 | 1.52 | 2.36 | 1.06 | 0 | 0.53 | 0.3 | 2.87
| Analytical results for Li in OREAS 550b (Certified Value 78 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 12 | 17 | 18 | 19 | 20 | 25 | 26 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [8 empty] | 0
1
| 1 | 1 | 83 | 86 | 72 | 65 | 81 | 70 | 83.608 | 80 | 74 | 79 | 94 | 80 | 89.1 | 75.7 | [8 empty] | 1
| 1 | 2 | 85 | 79 | 72 | 80 | 81 | 80 | 84.3949 | 80 | 74 | 78 | 75 | 80 | 71.9 | 79.2 | [8 empty] | #N/A
| 1 | 3 | 87 | 78 | 72 | 70 | 80 | 80 | 82.5811 | 80 | 73 | 77 | 83 | 70 | 74.5 | 73.3 | [8 empty] | 16
| 1 | 4 | 86 | 80 | 75 | 90 | 81 | 80 | 84.3094 | 80 | 73 | 78 | 82 | 70 | 82.4 | 71.6 | [8 empty] | 78.3646
| 1 | 5 | 85 | 79 | 74 | 75 | 80 | 70 | 82.8055 | 80 | 75 | 79 | 77 | 70 | 83.5 | 72.6 | [8 empty] | 116
| 1 | 6 | 83 | 79 | 72 | 80 | 80 | 80 | 82.7279 | 80 | 73 | 79 | 75 | 80 | 86.5 | 73.5
| Mean |  | 84.8333 | 80.1667 | 72.8333 | 76.6667 | 80.5 | 76.6667 | 83.4045 | 80 | 73.6667 | 78.3333 | 81 | 75 | 81.3167 | 74.3167
| Median |  | 85 | 79 | 72 | 77.5 | 80.5 | 80 | 83.2068 | 80 | 73.5 | 78.5 | 79.5 | 75 | 82.95 | 73.4
| Std Dev. |  | 1.6021 | 2.9269 | 1.3292 | 8.756 | 0.54772 | 5.164 | 0.81677 | 0 | 0.8165 | 0.8165 | 7.2388 | 5.4772 | 6.7606 | 2.7491
| Rel.Std.Dev. |  | 1.889% | 3.651% | 1.825% | 11.42% | 0.6804% | 6.736% | 0.9793% | 0% | 1.108% | 1.042% | 8.937% | 7.303% | 8.314% | 3.699%
| PDM3 |  | 8.255% | 2.3% | -7.058% | -2.167% | 2.725% | -2.167% | 6.431% | 2.087% | -5.995% | -0.0399% | 3.363% | -4.294% | 3.767% | -5.166%
| Z-Score (Absolute) |  | 1.53 | 0.27 | 1.71 | 0.67 | 0.36 | 0.67 | 1.14 | 0.22 | 1.48 | 0.22 | 0.49 | 1.12 | 0.58 | 1.31
| Analytical results for Lu in OREAS 550b (Certified Value 0.48 ppm) | [23 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 06 | 15 | 17 | 22 | 25 | 26 | [16 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [16 empty] | 2
3
| 1 | 1 | 0.43 | 0.44432 | 0.5 | < 0.5 | 0.49 | 0.5 | [16 empty] | 1
| 1 | 2 | 0.43 | 0.46044 | 0.5 | < 0.5 | 0.45 | 0.6 | [16 empty] | 8
| 1 | 3 | 0.45 | 0.45213 | 0.47 | < 0.5 | 0.46 | 0.5 | [16 empty] | 16
| 1 | 4 | 0.48 | 0.45622 | 0.49 | < 0.5 | 0.53 | 0.5 | [16 empty] | 0.47695
| 1 | 5 | 0.48 | 0.46017 | 0.51 | < 0.5 | 0.49 | 0.5 | [16 empty] | 117
| 1 | 6 | 0.52 | 0.44522 | 0.48 | < 0.5 | 0.43 | 0.5
| Mean |  | 0.465 | 0.45308 | 0.49167 |  | 0.475 | 0.51667
| Median |  | 0.465 | 0.45417 | 0.495 |  | 0.475 | 0.5
| Std Dev. |  | 0.035071 | 0.0071233 | 0.01472 |  | 0.035637 | 0.040825
| Rel.Std.Dev. |  | 7.542% | 1.572% | 2.994% |  | 7.503% | 7.902%
| PDM3 |  | -2.505% | -5.004% | 3.086% |  | -0.4088% | 8.327%
| Z-Score (Absolute) |  | 0.17 | 0.59 | 0.76 | 7.69 | 0.17 | 1.63
| Analytical results for Mg in OREAS 550b (Certified Value 4.35 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 4.38 | 4.342 | 4.42 | 4.18 | 3.4364 | 3.76 | 4.54 | 4.22 | 4.1323 | 4.312 | 4.282 | 4.318 | 4.209 | 4.36 | 4.4387 | 4.27 | 4.38 | 4.5249 | 4.31 | 3.12 | 3.71 | 4.29 | 1
| 1 | 2 | 4.48 | 4.384 | 4.33 | 4.11 | 3.437 | 3.69 | 4.53 | 4.2 | 4.3789 | 4.282 | 4.493 | 4.384 | 4.27 | 4.348 | 4.4955 | 4.37 | 4.38 | 4.5034 | 4.26 | 3.11 | 3.6 | 4.22 | #N/A
| 1 | 3 | 4.37 | 4.469 | 4.28 | 4.16 | 3.4396 | 3.77 | 4.52 | 4.16 | 4.6013 | 4.318 | 4.42 | 4.396 | 4.288 | 4.33 | 4.3546 | 4.35 | 4.31 | 4.6022 | 4.26 | 3.14 | 4.04 | 4.16 | 16
| 1 | 4 | 4.34 | 4.445 | 4.32 | 4.13 | 3.4373 | 3.99 | 4.53 | 4.2 | 4.4228 | 4.336 | 4.39 | 4.324 | 4.252 | 4.348 | 4.403 | 4.39 | 4.33 | 4.6502 | 4.19 | 3.13 | 4.46 | 4.16 | 4.3484
| 1 | 5 | 4.38 | 4.451 | 4.35 | 4.19 | 3.4303 | 3.95 | 4.51 | 4.23 | 4.3564 | 4.324 | 4.36 | 4.354 | 4.258 | 4.439 | 4.3565 | 4.35 | 4.59 | 4.5132 | 4.22 | 3.12 | 4.62 | 4.1 | 118
| 1 | 6 | 4.36 | 4.601 | 4.34 | 4.18 | 3.4392 | 3.92 | 4.51 | 4.18 | 4.296 | 4.372 | 4.402 | 4.451 | 4.288 | 4.39 | 4.3096 | 4.25 | 4.47 | 4.5798 | 4.31 | 3.11 | 4.15 | 4.35
| Mean |  | 4.385 | 4.4487 | 4.34 | 4.1583 | 3.4366 | 3.8467 | 4.5233 | 4.1983 | 4.3646 | 4.324 | 4.3912 | 4.3712 | 4.2608 | 4.3692 | 4.393 | 4.33 | 4.41 | 4.5623 | 4.2583 | 3.1217 | 4.0967 | 4.2133
| Median |  | 4.375 | 4.448 | 4.335 | 4.17 | 3.4371 | 3.845 | 4.525 | 4.2 | 4.3676 | 4.321 | 4.396 | 4.369 | 4.264 | 4.354 | 4.3797 | 4.35 | 4.38 | 4.5523 | 4.26 | 3.12 | 4.095 | 4.19
| Std Dev. |  | 0.048888 | 0.088541 | 0.046043 | 0.031885 | 0.0033506 | 0.12209 | 0.012111 | 0.025626 | 0.15385 | 0.029638 | 0.069583 | 0.050002 | 0.029438 | 0.03956 | 0.067033 | 0.056569 | 0.10412 | 0.058144 | 0.047924 | 0.01169 | 0.40203 | 0.09288
| Rel.Std.Dev. |  | 1.115% | 1.99% | 1.061% | 0.7668% | 0.0975% | 3.174% | 0.2677% | 0.6104% | 3.525% | 0.6854% | 1.585% | 1.144% | 0.6909% | 0.9054% | 1.526% | 1.306% | 2.361% | 1.274% | 1.125% | 0.3745% | 9.814% | 2.204%
| PDM3 |  | 0.8425% | 2.307% | -0.1924% | -4.37% | -20.97% | -11.54% | 4.024% | -3.45% | 0.3735% | -0.5604% | 0.9843% | 0.5243% | -2.013% | 0.4783% | 1.026% | -0.4224% | 1.417% | 4.919% | -2.071% | -28.21% | -5.788% | -3.105%
| Z-Score (Absolute) |  | 0.44 | 1.01 | 0.04 | 1.57 | 7.99 | 4.34 | 1.67 | 1.22 | 0.26 | 0.1 | 0.5 | 0.32 | 0.66 | 0.3 | 0.52 | 0.04 | 0.67 | 2.02 | 0.68 | 10.79 | 2.12 | 1.08
| Analytical results for Mn in OREAS 550b (Certified Value 0.02 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.0196 | 0.02 | 0.0195 | 0.019 | 0.020161 | 0.019 | 0.0221 | 0.018 | 0.01837 | 0.015 | 0.015 | 0.015 | 0.015 | 0.023 | 0.019069 | 0.0192 | 0.0193 | 0.0192 | 0.02 | 0.01816 | 0.02 | 0.0216 | 1
| 1 | 2 | 0.0193 | 0.02 | 0.0193 | 0.019 | 0.020567 | 0.0184 | 0.022 | 0.019 | 0.019381 | 0.015 | 0.023 | 0.023 | 0.023 | 0.023 | 0.019204 | 0.0195 | 0.0193 | 0.0176 | 0.02 | 0.01802 | 0.021 | 0.0214 | #N/A
| 1 | 3 | 0.0193 | 0.02 | 0.0193 | 0.019 | 0.020505 | 0.019 | 0.022 | 0.018 | 0.019089 | 0.023 | 0.015 | 0.023 | 0.015 | 0.023 | 0.018567 | 0.0194 | 0.0191 | 0.0193 | 0.02 | 0.01814 | 0.02 | 0.0207 | 16
| 1 | 4 | 0.0192 | 0.02 | 0.0192 | 0.019 | 0.020167 | 0.0195 | 0.022 | 0.019 | 0.018274 | 0.023 | 0.015 | 0.023 | 0.023 | 0.023 | 0.018199 | 0.0195 | 0.0191 | 0.0178 | 0.02 | 0.0181 | 0.019 | 0.0213 | 0.019522
| 1 | 5 | 0.0195 | 0.02 | 0.0206 | 0.019 | 0.020337 | 0.0204 | 0.022 | 0.019 | 0.018053 | 0.023 | 0.015 | 0.023 | 0.015 | 0.023 | 0.019192 | 0.0195 | 0.0202 | 0.0184 | 0.02 | 0.01799 | 0.019 | 0.021 | 119
| 1 | 6 | 0.0193 | 0.02 | 0.0207 | 0.019 | 0.020489 | 0.02 | 0.022 | 0.018 | 0.019362 | 0.023 | 0.015 | 0.023 | 0.015 | 0.023 | 0.018379 | 0.0189 | 0.0199 | 0.0183 | 0.02 | 0.01802 | 0.02 | 0.0226
| Mean |  | 0.019367 | 0.02 | 0.019767 | 0.019 | 0.020371 | 0.019383 | 0.022017 | 0.0185 | 0.018755 | 0.020333 | 0.016333 | 0.021667 | 0.017667 | 0.023 | 0.018768 | 0.019333 | 0.019483 | 0.018433 | 0.02 | 0.018072 | 0.019833 | 0.021433
| Median |  | 0.0193 | 0.02 | 0.0194 | 0.019 | 0.020413 | 0.01925 | 0.022 | 0.0185 | 0.01873 | 0.023 | 0.015 | 0.023 | 0.015 | 0.023 | 0.018818 | 0.01945 | 0.0193 | 0.01835 | 0.02 | 0.01806 | 0.02 | 0.02135
| Std Dev. |  | 0.00015055 | 0 | 0.00069186 | 0 | 0.00017739 | 0.00073326 | 4.0825e-05 | 0.00054772 | 0.00059043 | 0.0041312 | 0.003266 | 0.003266 | 0.0041312 | 3.8006e-18 | 0.00044167 | 0.00024221 | 0.00045789 | 0.00070048 | 0 | 7.111e-05 | 0.00075277 | 0.0006532
| Rel.Std.Dev. |  | 0.7774% | 0% | 3.5% | 0% | 0.8708% | 3.783% | 0.1854% | 2.961% | 3.148% | 20.32% | 20% | 15.07% | 23.38% | 1.652e-14% | 2.353% | 1.253% | 2.35% | 3.8% | 0% | 0.3935% | 3.795% | 3.048%
| PDM3 |  | -0.7956% | 2.449% | 1.253% | -2.674% | 4.349% | -0.7102% | 12.78% | -5.235% | -3.93% | 4.156% | -16.33% | 10.99% | -9.504% | 17.82% | -3.861% | -0.9663% | -0.1979% | -5.576% | 2.449% | -7.429% | 1.595% | 9.791%
| Z-Score (Absolute) |  | 0.06 | 0.47 | 0.28 | 0.36 | 0.78 | 0.04 | 2.16 | 0.78 | 0.57 | 1.02 | 2.22 | 2.1 | 1.14 | 3.18 | 0.56 | 0.08 | 0.04 | 0.84 | 0.47 | 1.14 | 0.33 | 1.67
| Analytical results for Mo in OREAS 550b (Certified Value 7.07 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [10 empty] | 2
2
| 1 | 1 | 7 | 8 | 7 | 6 | < 10 | 7.6369 | 6.2649 | 7 | 7 | 5.9 | 14 | 8 | [10 empty] | 1
| 1 | 2 | 7 | 8 | 7 | 6 | < 10 | 7.8078 | 6.6876 | 7 | 7 | 5.9 | 13 | 8 | [10 empty] | 9
| 1 | 3 | 8 | 8 | 7 | 6 | < 10 | 7.55 | 6.373 | 7 | 8 | 6 | 11 | 8 | [10 empty] | 16
| 1 | 4 | 8 | 7 | 7 | 6 | < 10 | 7.4725 | 5.9894 | 7 | 8 | 5.9 | 11 | 7 | [10 empty] | 7.0692
| 1 | 5 | 8 | 9 | 7 | 6 | < 10 | 7.1436 | 6.6935 | 7 | 7 | 5.9 | 16 | 8 | [10 empty] | 120
| 1 | 6 | 7 | 9 | 7 | 6 | < 10 | 7.0509 | 6.9826 | 7 | 7 | 5.9 | 10 | 8
| Mean |  | 7.5 | 8.1667 | 7 | 6 |  | 7.4436 | 6.4985 | 7 | 7.3333 | 5.9167 | 12.5 | 7.8333
| Median |  | 7.5 | 8 | 7 | 6 |  | 7.5113 | 6.5303 | 7 | 7 | 5.9 | 12 | 8
| Std Dev. |  | 0.54772 | 0.75277 | 0 | 0 |  | 0.29201 | 0.35711 | 0 | 0.5164 | 0.040825 | 2.2583 | 0.40825
| Rel.Std.Dev. |  | 7.303% | 9.218% | 0% | 0% |  | 3.923% | 5.495% | 0% | 7.042% | 0.69% | 18.07% | 5.212%
| PDM3 |  | 6.094% | 15.52% | -0.9791% | -15.12% |  | 5.297% | -8.073% | -0.9791% | 3.736% | -16.3% | 76.82% | 10.81%
| Z-Score (Absolute) |  | 0.34 | 1.01 | 0.17 | 1.18 | 2.19 | 0.28 | 0.68 | 0.17 | 0.17 | 1.26 | 5.39 | 0.67
| Analytical results for Na in OREAS 550b (Indicative Value 0.039 wt.%) | [23 empty] | Indicative
Na | Round | Replicate | Lab | Lab | [20 empty] | 1
| No | No | 02 | 22 | [20 empty] | wt.%
|  |  | BF*OES | PF*OES | [20 empty] | 3
3
| 1 | 1 | 0.037 | 0.01 | [20 empty] | 1
| 1 | 2 | 0.111 | 0.01 | [20 empty] | 21
| 1 | 3 | 0.03 | 0.02 | [20 empty] | 16
| 1 | 4 | 0.052 | 0.01 | [20 empty] | 0.039216
| 1 | 5 | 0.111 | 0.01 | [20 empty] | 27
| 1 | 6 | 0.059 | 0.01
| Mean |  | 0.066667 | 0.011667
| Median |  | 0.0555 | 0.01
| Std Dev. |  | 0.035859 | 0.0040825
| Rel.Std.Dev. |  | 53.79% | 34.99%
| PDM3 |  | 70% | -70.25%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 550b (Certified Value 19.5 ppm) | [23 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 06 | 09 | 12 | 14 | 15 | 17 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [12 empty] | 1
2
| 1 | 1 | 19 | 19 | 20.8039 | 22 | < 50 | 18.3663 | 18 | 18.5 | 24 | 21 | [12 empty] | 1
| 1 | 2 | 20.4 | 20 | 21.2912 | 21 | < 50 | 18.2981 | 18 | 18.2 | 27 | 21 | [12 empty] | 25
| 1 | 3 | 17.2 | 18 | 20.797 | 22 | < 50 | 18.7461 | 18 | 17.7 | 25 | 21 | [12 empty] | 16
| 1 | 4 | 18.7 | 20 | 19.2667 | 21 | < 50 | 19.0951 | 19 | 18.1 | 26 | 21 | [12 empty] | 19.4534
| 1 | 5 | 18.8 | 19 | 18.8393 | 21 | < 50 | 18.8839 | 18 | 18.8 | 22 | 21 | [12 empty] | 121
| 1 | 6 | 19.8 | 20 | 18.6161 | 21 | < 50 | 18.1599 | 18 | 18.4 | 23 | 21
| Mean |  | 18.9833 | 19.3333 | 19.9357 | 21.3333 |  | 18.5916 | 18.1667 | 18.2833 | 24.5 | 21
| Median |  | 18.9 | 19.5 | 20.0319 | 21 |  | 18.5562 | 18 | 18.3 | 24.5 | 21
| Std Dev. |  | 1.0926 | 0.8165 | 1.1597 | 0.5164 |  | 0.37043 | 0.40825 | 0.37639 | 1.8708 | 0
| Rel.Std.Dev. |  | 5.755% | 4.223% | 5.817% | 2.421% |  | 1.992% | 2.247% | 2.059% | 7.636% | 0%
| PDM3 |  | -2.416% | -0.6172% | 2.479% | 9.664% |  | -4.43% | -6.614% | -6.015% | 25.94% | 7.95%
| Z-Score (Absolute) |  | 0.32 | 0.15 | 0.15 | 0.84 | 2.66 | 0.52 | 0.73 | 0.67 | 2.42 | 0.68
| Analytical results for Nd in OREAS 550b (Certified Value 53 ppm) | [23 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 0
1
| 1 | 1 | 51.8 | 52.6 | 53.2181 | 51.6 | 50 | 50.45 | 51.4 | 53.9 | [14 empty] | 1
| 1 | 2 | 52.2 | 53.5 | 53.2239 | 52.6 | 50 | 50.89 | 47.2 | 56.3 | [14 empty] | 12
| 1 | 3 | 51.7 | 51.9 | 50.6679 | 52.2 | 51 | 51.26 | 51.3 | 55 | [14 empty] | 16
| 1 | 4 | 57 | 59.8 | 51.4916 | 53.1 | 51 | 49.94 | 51.6 | 55.2 | [14 empty] | 52.6152
| 1 | 5 | 55 | 55.4 | 52.8189 | 52.4 | 50 | 48.56 | 51.7 | 54.9 | [14 empty] | 122
| 1 | 6 | 55.8 | 58.3 | 54.2292 | 53.1 | 51 | 50.66 | 53.9 | 53.4
| Mean |  | 53.9167 | 55.25 | 52.6083 | 52.5 | 50.5 | 50.2933 | 51.1833 | 54.7833
| Median |  | 53.6 | 54.45 | 53.0185 | 52.5 | 50.5 | 50.555 | 51.5 | 54.95
| Std Dev. |  | 2.3051 | 3.2042 | 1.2988 | 0.57271 | 0.54772 | 0.95695 | 2.1794 | 1.0226
| Rel.Std.Dev. |  | 4.275% | 5.799% | 2.469% | 1.091% | 1.085% | 1.903% | 4.258% | 1.867%
| PDM3 |  | 2.474% | 5.008% | -0.01316% | -0.219% | -4.02% | -4.413% | -2.721% | 4.121%
| Z-Score (Absolute) |  | 0.54 | 1.06 | 0.02 | 0.02 | 0.81 | 0.89 | 0.54 | 0.88
| Analytical results for Ni in OREAS 550b (Certified Value 70 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 22 | 25 | 26 |  |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 0
0
| 1 | 1 | 63 | 80 | 80 | 80 | 72 | 70 | 60.096 | 50 | 60 | 70 | 120 | 100 | 67 | 62 | 50 | 69.41 | 56 | 70 |  |  |  |  | 1
| 1 | 2 | 71 | 70 | 70 | 80 | 72 | 70 | 63.1858 | 70 | 60 | 70 | 110 | 80 | 69 | 59 | < 50 | 69.34 | 54 | 85 |  |  |  |  | 28
| 1 | 3 | 62 | 70 | 80 | 80 | 72 | 60 | 65.9466 | 80 | 60 | 70 | 100 | 80 | 70 | 76 | < 50 | 69.37 | 62 | 74 |  |  |  |  | 16
| 1 | 4 | 70 | 70 | 80 | 70 | 72 | 60 | 63.2842 | 80 | 60 | 80 | 310 | 80 | 69 | 60 | < 50 | 69.39 | 63 | 75 |  |  |  |  | 69.5639
| 1 | 5 | 52 | 70 | 90 | 80 | 72 | 70 | 62.5028 | 80 | 70 | 90 | 140 | 80 | 70 | 72 | 60 | 69.42 | 66 | 75 |  |  |  |  | 123
| 1 | 6 | 58 | 80 | 80 | 80 | 72 | 60 | 67.1951 | 70 | 70 | 70 | 120 | 100 | 66 | 79 | 50 | 69.38 | 59 | 91
| Mean |  | 62.6667 | 73.3333 | 80 | 78.3333 | 72 | 65 | 63.7018 | 71.6667 | 63.3333 | 75 | 150 | 86.6667 | 68.5 | 68 | 53.3333 | 69.385 | 60 | 78.3333
| Median |  | 62.5 | 70 | 80 | 80 | 72 | 65 | 63.235 | 75 | 60 | 70 | 120 | 80 | 69 | 67 | 50 | 69.385 | 60.5 | 75
| Std Dev. |  | 7.2019 | 5.164 | 6.3246 | 4.0825 | 0 | 5.4772 | 2.5346 | 11.6905 | 5.164 | 8.3666 | 79.4984 | 10.328 | 1.6432 | 8.7407 | 5.7735 | 0.02881 | 4.5166 | 7.9415
| Rel.Std.Dev. |  | 11.49% | 7.042% | 7.906% | 5.212% | 0% | 8.427% | 3.979% | 16.31% | 8.154% | 11.16% | 53% | 11.92% | 2.399% | 12.85% | 10.83% | 0.04152% | 7.528% | 10.14%
| PDM3 |  | -9.915% | 5.419% | 15% | 12.61% | 3.502% | -6.561% | -8.427% | 3.023% | -8.957% | 7.815% | 115.6% | 24.59% | -1.529% | -2.248% | -23.33% | -0.2572% | -13.75% | 12.61%
| Z-Score (Absolute) |  | 0.76 | 0.27 | 0.91 | 0.75 | 0.14 | 0.53 | 0.66 | 0.11 | 0.69 | 0.43 | 7.65 | 1.55 | 0.19 | 0.24 | 3.02 | 0.11 | 1.01 | 0.75
| Analytical results for P in OREAS 550b (Certified Value 0.052 wt.%) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 06 | 07 | 08 | 17 | 19 | 22 | 25 | 26 | [12 empty] | wt.%
|  |  | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.048 | 0.05 | 0.05 | 0.0737 | 0.04 | 0.05 | 0.0569 | 0.05 | 0.04 | 0.07 | [12 empty] | 1
| 1 | 2 | 0.048 | 0.05 | 0.06 | 0.0738 | 0.05 | 0.05 | 0.0522 | 0.05 | 0.06 | 0.05 | [12 empty] | #N/A
| 1 | 3 | 0.052 | 0.05 | 0.05 | 0.0738 | 0.04 | 0.05 | 0.0589 | 0.05 | 0.05 | 0.04 | [12 empty] | 16
| 1 | 4 | 0.048 | 0.05 | 0.05 | 0.0738 | 0.04 | 0.05 | 0.058 | 0.06 | 0.06 | 0.06 | [12 empty] | 0.05178
| 1 | 5 | 0.052 | 0.05 | 0.05 | 0.0739 | 0.05 | 0.05 | 0.0572 | 0.05 | 0.04 | 0.08 | [12 empty] | 124
| 1 | 6 | 0.052 | 0.05 | 0.05 | 0.0738 | 0.04 | 0.05 | 0.0554 | 0.05 | 0.06 | 0.08
| Mean |  | 0.05 | 0.05 | 0.051667 | 0.0738 | 0.043333 | 0.05 | 0.056433 | 0.051667 | 0.051667 | 0.063333
| Median |  | 0.05 | 0.05 | 0.05 | 0.0738 | 0.04 | 0.05 | 0.05705 | 0.05 | 0.055 | 0.065
| Std Dev. |  | 0.0021909 | 7.6012e-18 | 0.0040825 | 6.3246e-05 | 0.005164 | 7.6012e-18 | 0.0023805 | 0.0040825 | 0.0098319 | 0.01633
| Rel.Std.Dev. |  | 4.382% | 1.52e-14% | 7.902% | 0.0857% | 11.92% | 1.52e-14% | 4.218% | 7.902% | 19.03% | 25.78%
| PDM3 |  | -3.438% | -3.438% | -0.2198% | 42.52% | -16.31% | -3.438% | 8.986% | -0.2198% | -0.2198% | 22.31%
| Z-Score (Absolute) |  | 0.6 | 0.67 | 0 | 8.95 | 3.37 | 0.67 | 1.93 | 0 | 0 | 4.72
| Analytical results for Pb in OREAS 550b (Certified Value 15 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 22 | 25 | 26 | [6 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | [6 empty] | 1
2
| 1 | 1 | < 100 | 16 | < 20 | 50 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 14 |  | < 300 | 14.8 | 17 | 15 | [6 empty] | 1
| 1 | 2 | < 100 | 13.5 | < 20 | 30 | < 20 | < 100 | < 100 | < 100 | < 100 | 200 | 15 |  | < 300 | 14.9 | 17 | 15 | [6 empty] | 13
| 1 | 3 | < 100 | 11.5 | < 20 | 40 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 16 |  | < 300 | 14.9 | 16 | 15 | [6 empty] | 16
| 1 | 4 | < 100 | 14.8 | < 20 | 50 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 15 | 33 | < 300 | 14.9 | 17 | 15 | [6 empty] | 14.9633
| 1 | 5 | < 100 | 11.5 | < 20 | 40 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 15 |  | < 300 | 14.8 | 16 | 15 | [6 empty] | 125
| 1 | 6 | < 100 | 13.4 | < 20 | 40 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 15 |  | < 300 | 14.9 | 16 | 15
| Mean |  |  | 13.45 |  | 41.6667 | [5 empty] | 200 | 15 | 33 |  | 14.8667 | 16.5 | 15
| Median |  |  | 13.45 |  | 40 | [5 empty] | 200 | 15 | 33 |  | 14.9 | 16.5 | 15
| Std Dev. |  |  | 1.7852 |  | 7.5277 | [6 empty] | 0.63246 |  |  | 0.05164 | 0.54772 | 0
| Rel.Std.Dev. |  |  | 13.27% |  | 18.07% | [6 empty] | 4.216% |  |  | 0.3474% | 3.32% | 0%
| PDM3 |  |  | -10.11% |  | 178.5% | [5 empty] | 1237% | 0.245% | 120.5% |  | -0.646% | 10.27% | 0.245%
| Z-Score (Absolute) |  | 0.4 | 0.75 | 0.85 | 0.14 | 0.85 | 0.4 | 0.4 | 0.4 | 0.4 | 1.18 | 0.7 | 0.14 | 3.52 | 0.7 | 0.65 | 0.7
| Analytical results for Pr in OREAS 550b (Certified Value 14.4 ppm) | [23 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 22 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [15 empty] | 1
2
| 1 | 1 | 13.5 | 15 | 13.9907 | 14.64 | 14.3 | 13 | 14.91 | [15 empty] | 1
| 1 | 2 | 14.3 | 14.81 | 14.3077 | 14.77 | 14.3 | 11.9 | 15.61 | [15 empty] | 14
| 1 | 3 | 11.9 | 14.88 | 14.4402 | 14.89 | 14.4 | 12.7 | 15.24 | [15 empty] | 16
| 1 | 4 | 15 | 16.55 | 14.3783 | 14.71 | 14.1 | 13.6 | 15.18 | [15 empty] | 14.4294
| 1 | 5 | 14 | 15.73 | 14.7982 | 14.86 | 13.9 | 13.3 | 15.52 | [15 empty] | 126
| 1 | 6 | 13.8 | 16.26 | 14.6794 | 15.14 | 14.4 | 13.5 | 14.84
| Mean |  | 13.75 | 15.5383 | 14.4324 | 14.835 | 14.2333 | 13 | 15.2167
| Median |  | 13.9 | 15.365 | 14.4092 | 14.815 | 14.3 | 13.15 | 15.21
| Std Dev. |  | 1.0407 | 0.75295 | 0.28581 | 0.17581 | 0.19664 | 0.63246 | 0.3113
| Rel.Std.Dev. |  | 7.569% | 4.846% | 1.98% | 1.185% | 1.382% | 4.865% | 2.046%
| PDM3 |  | -4.708% | 7.685% | 0.02095% | 2.811% | -1.359% | -9.906% | 5.456%
| Z-Score (Absolute) |  | 0.67 | 1.09 | 0 | 0.4 | 0.2 | 1.42 | 0.77
| Analytical results for Rb in OREAS 550b (Certified Value 88 ppm) | [23 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 14 | 15 | 17 | 18 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 0
1
| 1 | 1 | 86.3 | 88.5 | 116 | 102 | 83.0898 | 96.4 | 90 | 90.2115 | 87.7 | 86.2 | 92 | 83 | [10 empty] | 1
| 1 | 2 | 87.5 | 85.4 | 111 | 99 | 86.3519 | 92.3 | 90 | 91.9183 | 87.4 | 86.2 | 95 | 87.5 | [10 empty] | 15
| 1 | 3 | 86.1 | 92.7 | 115 | 95 | 85.6779 | 93.3 | 80 | 90.1338 | 86.2 | 87.7 | 84 | 83.9 | [10 empty] | 16
| 1 | 4 | 84 | 82.4 | 115 | 103 | 82.5966 | 94.8 | 90 | 91.0777 | 87 | 87.3 | 90 | 82.4 | [10 empty] | 87.7846
| 1 | 5 | 85.5 | 89.7 | 116 | 98 | 81.7136 | 91.3 | 90 | 90.2744 | 85.3 | 87.7 | 88 | 88.6 | [10 empty] | 127
| 1 | 6 | 87.4 | 90.8 | 117 | 102 | 84.4675 | 93 | 90 | 89.1652 | 85.4 | 86.3 | 83 | 85.2
| Mean |  | 86.1333 | 88.25 | 115 | 99.8333 | 83.9829 | 93.5167 | 88.3333 | 90.4635 | 86.5 | 86.9 | 88.6667 | 85.1
| Median |  | 86.2 | 89.1 | 115.5 | 100.5 | 83.7787 | 93.15 | 90 | 90.243 | 86.6 | 86.8 | 89 | 84.55
| Std Dev. |  | 1.3003 | 3.7634 | 2.0976 | 3.0605 | 1.8215 | 1.8258 | 4.0825 | 0.93672 | 1.0237 | 0.74565 | 4.6332 | 2.4964
| Rel.Std.Dev. |  | 1.51% | 4.264% | 1.824% | 3.066% | 2.169% | 1.952% | 4.622% | 1.035% | 1.183% | 0.8581% | 5.225% | 2.933%
| PDM3 |  | -1.881% | 0.5301% | 31% | 13.73% | -4.331% | 6.53% | 0.625% | 3.052% | -1.463% | -1.008% | 1.005% | -3.058%
| Z-Score (Absolute) |  | 0.67 | 0.01 | 8.32 | 3.59 | 1.34 | 1.63 | 0.01 | 0.68 | 0.56 | 0.43 | 0.12 | 0.99
| Analytical results for Re in OREAS 550b (Certified Value < 0.1 ppm) | [23 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 04 | 06 | 09 | 25 | 26 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [17 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 1
| 1 | 2 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 16
| 1 | 3 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 16
| 1 | 4 | < 0.1 | 0.06 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 128
| 1 | 6 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05
| Mean |  |  | 0.06
| Median |  |  | 0.06
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 550b (Certified Value 1.18 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 19 | 20 | 22 | 25 | 26 |  |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 1.19 | 1.17 | 1.18 | 1.0914 | 1.05 | 1.17 | 1.01 | 1.2207 | 1.18 | 1.16 | 1.19 | 1.22 | 1.21 | 1.16 | 1.2251 | 1.16 | 1.42 | 1.23 | 1.15 |  |  |  | 1
| 1 | 2 | 1.21 | 1.15 | 1.2 | 1.042 | 1.02 | 1.17 | 0.94 | 1.2619 | 1.19 | 1.21 | 1.19 | 1.26 | 1.23 | 1.21 | 1.1833 | 1.14 | 1.39 | 1.26 | 1.14 |  |  |  | 30
| 1 | 3 | 1.21 | 1.11 | 1.17 | 1.0751 | 1.05 | 1.18 | 0.95 | 1.2594 | 1.2 | 1.19 | 1.21 | 1.22 | 1.2 | 1.14 | 1.2241 | 1.15 | 1.42 | 1.32 | 1.14 |  |  |  | 16
| 1 | 4 | 1.26 | 1.12 | 1.17 | 1.0567 | 1.13 | 1.17 | 0.98 | 1.2705 | 1.2 | 1.18 | 1.18 | 1.21 | 1.2 | 1.15 | 1.3107 | 1.13 | 1.41 | 1.25 | 1.13 |  |  |  | 1.179
| 1 | 5 | 1.19 | 1.1 | 1.2 | 1.0583 | 1.04 | 1.18 | 1.02 | 1.2441 | 1.2 | 1.17 | 1.17 | 1.24 | 1.24 | 1.22 | 1.1574 | 1.14 | 1.37 | 1.28 | 1.13 |  |  |  | 129
| 1 | 6 | 1.21 | 1.11 | 1.19 | 1.0652 | 1.11 | 1.18 | 0.93 | 1.2234 | 1.19 | 1.18 | 1.23 | 1.19 | 1.24 | 1.18 | 1.2676 | 1.15 | 1.38 | 1.3 | 1.15
| Mean |  | 1.2117 | 1.1267 | 1.185 | 1.0648 | 1.0667 | 1.175 | 0.97167 | 1.2467 | 1.1933 | 1.1817 | 1.195 | 1.2233 | 1.22 | 1.1767 | 1.228 | 1.145 | 1.3983 | 1.2733 | 1.14
| Median |  | 1.21 | 1.115 | 1.185 | 1.0618 | 1.05 | 1.175 | 0.965 | 1.2518 | 1.195 | 1.18 | 1.19 | 1.22 | 1.22 | 1.17 | 1.2246 | 1.145 | 1.4 | 1.27 | 1.14
| Std Dev. |  | 0.025626 | 0.027325 | 0.013784 | 0.01698 | 0.043205 | 0.0054772 | 0.037639 | 0.020895 | 0.008165 | 0.017224 | 0.021679 | 0.024221 | 0.018974 | 0.03266 | 0.055522 | 0.010488 | 0.02137 | 0.033267 | 0.0089443
| Rel.Std.Dev. |  | 2.115% | 2.425% | 1.163% | 1.595% | 4.05% | 0.4661% | 3.874% | 1.676% | 0.6842% | 1.458% | 1.814% | 1.98% | 1.555% | 2.776% | 4.521% | 0.916% | 1.528% | 2.613% | 0.7846%
| PDM3 |  | 2.77% | -4.44% | 0.5079% | -9.686% | -9.529% | -0.3402% | -17.59% | 5.74% | 1.215% | 0.2252% | 1.356% | 3.759% | 3.477% | -0.1989% | 4.158% | -2.885% | 18.6% | 8% | -3.309%
| Z-Score (Absolute) |  | 0.45 | 0.98 | 0 | 2.03 | 1.99 | 0.17 | 3.6 | 1.04 | 0.14 | 0.06 | 0.17 | 0.65 | 0.59 | 0.14 | 0.73 | 0.67 | 3.6 | 1.49 | 0.76
| Analytical results for Sb in OREAS 550b (Certified Value 2.67 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 06 | 08 | 09 | 17 | 18 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [12 empty] | 2
3
| 1 | 1 | 2.9 | 3 | 2 | < 50 | 2.2084 | 3 | 2.7 | < 5 | 2.7 | 3.1 | [12 empty] | 1
| 1 | 2 | 2.5 | 4 | 1 | < 50 | 2.6316 | 3.4 | 2.5 | < 5 | 2.7 | 3.3 | [12 empty] | 17
| 1 | 3 | 2.8 | 4 | 1 | < 50 | 2.5851 | 2.7 | 2.4 | < 5 | 2.4 | 3.3 | [12 empty] | 16
| 1 | 4 | 3.1 | 4 | 1 | < 50 | 2.2663 | 2.9 | 2.5 | < 5 | 2.6 | 3.7 | [12 empty] | 2.6653
| 1 | 5 | 2.7 | 4 | < 1 | < 50 | 2.6938 | 2.9 | 2.5 | < 5 | 2.7 | 3.6 | [12 empty] | 130
| 1 | 6 | 2.8 | 4 | 2 | < 50 | 2.7133 | 2.8 | 2.7 | < 5 | 2.5 | 3.6
| Mean |  | 2.8 | 3.8333 | 1.4 |  | 2.5164 | 2.95 | 2.55 |  | 2.6 | 3.4333
| Median |  | 2.8 | 4 | 1 |  | 2.6083 | 2.9 | 2.5 |  | 2.65 | 3.45
| Std Dev. |  | 0.2 | 0.40825 | 0.54772 |  | 0.22165 | 0.2429 | 0.12247 |  | 0.12649 | 0.23381
| Rel.Std.Dev. |  | 7.143% | 10.65% | 39.12% |  | 8.808% | 8.234% | 4.803% |  | 4.865% | 6.81%
| PDM3 |  | 5.055% | 43.82% | -47.47% |  | -5.586% | 10.68% | -4.325% |  | -2.449% | 28.82%
| Z-Score (Absolute) |  | 0.35 | NA | NA | 78.4 | 0.65 | 0.88 | 0.53 | 0.7 | 0.35 | 2.58
Results from laboratories 3 and 6 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 550b (Certified Value 11.1 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 04 | 07 | 08 | 09 | 17 | 18 | 19 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | [10 empty] | 1
2
| 1 | 1 |  | 13 | < 20 | 11 | 11 | < 20 | 11 | 10 | 11 | 11 | 10 | 11 | [10 empty] | 1
| 1 | 2 |  | 11 | < 20 | 11 | 12 | < 20 | 11 |  | 11 | 10.8 | 13 | 12 | [10 empty] | 18
| 1 | 3 |  | 11 | < 20 | 11 | 11 | < 20 | 11 | 10 | 11 | 10.9 | 13 | 12 | [10 empty] | 16
| 1 | 4 | 11 | 13 | < 20 | 11 | 12 | < 20 | 11 |  | 11 | 10.9 | 11 | 11 | [10 empty] | 11.12
| 1 | 5 | 10 | 11 | < 20 | 11 | 12 | < 20 | 11 | 11 | 11 | 11.6 | 13 | 11 | [10 empty] | 131
| 1 | 6 |  | 10 | < 20 | 11 | 11 | < 20 | 11 | 10 | 11 | 11.5 | 12 | 11
| Mean |  | 10.5 | 11.5 |  | 11 | 11.5 |  | 11 | 10.25 | 11 | 11.1167 | 12 | 11.3333
| Median |  | 10.5 | 11 |  | 11 | 11.5 |  | 11 | 10 | 11 | 10.95 | 12.5 | 11
| Std Dev. |  | 0.70711 | 1.2247 |  | 0 | 0.54772 |  | 0 | 0.5 | 0 | 0.34303 | 1.2649 | 0.5164
| Rel.Std.Dev. |  | 6.734% | 10.65% |  | 0% | 4.763% |  | 0% | 4.878% | 0% | 3.086% | 10.54% | 4.556%
| PDM3 |  | -5.576% | 3.417% |  | -1.079% | 3.417% |  | -1.079% | -7.824% | -1.079% | -0.02998% | 7.914% | 1.918%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 1.35 | 0 | 0.67 | 1.35 | 0 | 1.01 | 0 | 0.16 | 1.35 | 0.45
| Analytical results for Se in OREAS 550b (Indicative Value 7.58 ppm) | [23 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 03 | 04 | 09 | 25 | 26 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [17 empty] | 2
2
| 1 | 1 | 13 | < 20 | < 20 | 2 | < 50 | [17 empty] | 1
| 1 | 2 | 13 | < 20 | < 20 | 2 | < 50 | [17 empty] | 22
| 1 | 3 | 12 | < 20 | < 20 | 2 | < 50 | [17 empty] | 16
| 1 | 4 | 15 | < 20 | < 20 | 3 | < 50 | [17 empty] | 7.5833
| 1 | 5 | 13 | < 20 | < 20 | 2 | < 50 | [17 empty] | 28
| 1 | 6 | 12 | < 20 | < 20 | 2 | < 50
| Mean |  | 13 |  |  | 2.1667
| Median |  | 13 |  |  | 2
| Std Dev. |  | 1.0954 |  |  | 0.40825
| Rel.Std.Dev. |  | 8.427% |  |  | 18.84%
| PDM3 |  | 71.43% |  |  | -71.43%
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 1.76 | 3.37
| Analytical results for Si in OREAS 550b (Certified Value 27.28 wt.%) | [23 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 25 | 26 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 2
3
| 1 | 1 | 27.5 | 27.1 | 28.3 | 26 | 26.5903 | 25.9 | 27.32 | 27.79 | 27.1024 | 26.6 | 26.6 | 27.39 | 28.42 | 27.81 | 26 | 24.97 | 29.2 | [5 empty] | 1
| 1 | 2 | 28.1 | 26 | 27.6 | 26.5 | 26.5081 | 25.9 | 27.34 | 27.58 | 27.2617 | 26.6 | 28 | 27.72 | 28.7 | 27.91 | 26.8 | 26.65 | 28.9 | [5 empty] | #N/A
| 1 | 3 | 27.5 | 26.7 | 26.1 | 26 | 26.4707 | 25.4 | 27.32 | 28.14 | 27.3899 | 26.69 | 27.49 | 27.81 | 28.98 | 27.58 | 26.5 | 25.08 | 28.5 | [5 empty] | 16
| 1 | 4 | 27.6 | 25.9 | 26.5 | 25.9 | 26.4531 | 27.6 | 27.38 | 27.4 | 27.1152 | 26.88 | 27.39 | 27.39 | 28.61 | 27.81 | 26.9 | 23.83 | 28.4 | [5 empty] | 27.2817
| 1 | 5 | 27.6 | 26.9 | 26.9 | 26 | 26.4831 | 26.1 | 27.36 | 28.35 | 27.4719 | 26.88 | 27.3 | 27.49 | 28.61 | 28.19 | 26.8 | 23.4 | 28.2 | [5 empty] | 132
| 1 | 6 | 27.4 | 27.5 | 26.5 | 26.5 | 26.5726 | 27.5 | 27.35 | 27.47 | 27.1024 | 27.02 | 27.58 | 28.19 | 28.98 | 28.19 | 26.3 | 24.14 | 29.4
| Mean |  | 27.6167 | 26.6833 | 26.9833 | 26.15 | 26.513 | 26.4 | 27.345 | 27.7883 | 27.2406 | 26.7783 | 27.3933 | 27.665 | 28.7167 | 27.915 | 26.55 | 24.6783 | 28.7667
| Median |  | 27.55 | 26.8 | 26.7 | 26 | 26.4956 | 26 | 27.345 | 27.685 | 27.1884 | 26.785 | 27.44 | 27.605 | 28.655 | 27.86 | 26.65 | 24.555 | 28.7
| Std Dev. |  | 0.24833 | 0.62743 | 0.82077 | 0.27386 | 0.056268 | 0.92087 | 0.023452 | 0.38332 | 0.16135 | 0.17348 | 0.45833 | 0.31034 | 0.22349 | 0.23898 | 0.35071 | 1.1643 | 0.4761
| Rel.Std.Dev. |  | 0.8992% | 2.351% | 3.042% | 1.047% | 0.2122% | 3.488% | 0.08576% | 1.379% | 0.5923% | 0.6479% | 1.673% | 1.122% | 0.7783% | 0.8561% | 1.321% | 4.718% | 1.655%
| PDM3 |  | 1.228% | -2.193% | -1.094% | -4.148% | -2.818% | -3.232% | 0.2319% | 1.857% | -0.1509% | -1.845% | 0.4091% | 1.405% | 5.26% | 2.321% | -2.682% | -9.543% | 5.443%
| Z-Score (Absolute) |  | 0.46 | 0.67 | 0.31 | 1.32 | 0.88 | 1.02 | 0.13 | 0.66 | 0 | 0.56 | 0.18 | 0.51 | 1.79 | 0.82 | 0.84 | 3.1 | 1.85
| Analytical results for Sm in OREAS 550b (Certified Value 7.7 ppm) | [23 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 6.9 | 7.1 | 8.3 | 7.823 | 7.7 | 7.6 | 7.8 | 8 | [14 empty] | 1
| 1 | 2 | 8.5 | 7.1 | 7.3 | 7.5908 | 7.7 | 7.4 | 7 | 8.2 | [14 empty] | 5
| 1 | 3 | 6.7 | 7.1 | 7.4 | 7.9272 | 8.1 | 7.2 | 7.5 | 8.1 | [14 empty] | 16
| 1 | 4 | 7.9 | 7.3 | 8.1 | 7.7304 | 8.2 | 7.6 | 7.7 | 8.2 | [14 empty] | 7.6994
| 1 | 5 | 7.7 | 7.4 | 8 | 7.7752 | 7.8 | 7.6 | 7.9 | 8.2 | [14 empty] | 133
| 1 | 6 | 7.3 | 7.3 | 8.4 | 7.827 | 8.2 | 7.5 | 7.8 | 8.1
| Mean |  | 7.5 | 7.2167 | 7.9167 | 7.7789 | 7.95 | 7.4833 | 7.6167 | 8.1333
| Median |  | 7.5 | 7.2 | 8.05 | 7.7991 | 7.95 | 7.55 | 7.75 | 8.15
| Std Dev. |  | 0.66933 | 0.13292 | 0.46224 | 0.11313 | 0.2429 | 0.16021 | 0.33116 | 0.08165
| Rel.Std.Dev. |  | 8.924% | 1.842% | 5.839% | 1.454% | 3.055% | 2.141% | 4.348% | 1.004%
| PDM3 |  | -2.59% | -6.27% | 2.821% | 1.032% | 3.254% | -2.807% | -1.075% | 5.635%
| Z-Score (Absolute) |  | 0.62 | 1.5 | 0.68 | 0.25 | 0.78 | 0.67 | 0.25 | 1.36
| Analytical results for Sn in OREAS 550b (Certified Value 3.02 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 06 | 08 | 09 | 12 | 14 | 15 | 17 | 20 | 22 | 25 | 26 | [9 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | [9 empty] | 2
2
| 1 | 1 | < 50 | 2.9 | < 10 | < 50 | < 100 | < 5 | < 5 | 2.8124 | 3 | < 50 | < 5 | 3.5 | < 10 | [9 empty] | 1
| 1 | 2 | < 50 | 4.3 | < 10 | < 50 | < 100 | < 5 | < 5 | 2.7056 | 3 | 70 | < 5 | 3.2 | < 10 | [9 empty] | 6
| 1 | 3 | < 50 | 3.9 | < 10 | < 50 | < 100 | < 5 | < 5 | 2.6726 | 3 | 60 | < 5 | 2.9 | < 10 | [9 empty] | 16
| 1 | 4 | < 50 | 3.3 | < 10 | < 50 | < 100 | < 5 | < 5 | 2.6072 | 3 | 120 | < 5 | 3.1 | < 10 | [9 empty] | 3.0195
| 1 | 5 | < 50 | 4.4 | < 10 | < 50 | < 100 | 5 | < 5 | 2.5821 | 3 | < 50 | < 5 | 2.8 | < 10 | [9 empty] | 134
| 1 | 6 | < 50 | 4.6 | < 10 | < 50 | < 100 | < 5 | < 5 | 2.604 | 3 | < 50 | < 5 | 2.7 | < 10
| Mean |  |  | 3.9 |  |  |  | 5 |  | 2.664 | 3 | 83.3333 |  | 3.0333
| Median |  |  | 4.1 |  |  |  | 5 |  | 2.6399 | 3 | 70 |  | 3
| Std Dev. |  |  | 0.67231 | [5 empty] | 0.086442 | 0 | 32.1455 |  | 0.29439
| Rel.Std.Dev. |  |  | 17.24% | [5 empty] | 3.245% | 0% | 38.57% |  | 9.705%
| PDM3 |  |  | 29.16% |  |  |  | 65.59% |  | -11.77% | -0.6446% | 2660% |  | 0.4593%
| Z-Score (Absolute) |  | 11.51 | 0 | 0.6 | 11.51 | 25.15 | 0.54 | 0.76 | 0.67 | 0.49 | 27.42 | 0.76 | 0.47 | 0.6
| Analytical results for Sr in OREAS 550b (Certified Value 78 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [7 empty] | 0
1
| 1 | 1 | 80 | 60 | 85 | 80 | 85 | 53 | 70 | 80.4306 | 72.2669 | 74 | 76 | 79 | 27.2 | 81 | 75.1 | [7 empty] | 1
| 1 | 2 | 82 | 80 | 83 | 80 | 85 | 53 | 80 | 81.5956 | 79.7814 | 78 | 76 | 77 | 27 | 83 | 82.3 | [7 empty] | #N/A
| 1 | 3 | 81 | 80 | 84 | 80 | 85 | 52 | 70 | 78.3948 | 75.6825 | 77 | 76 | 79 | 27.1 | 71 | 77.2 | [7 empty] | 16
| 1 | 4 | 78 | 70 | 83 | 80 | 90 | 53 | 70 | 81.788 | 77.4964 | 77 | 73 | 78 | 27.1 | 82 | 75.1 | [7 empty] | 78.2221
| 1 | 5 | 81 | 80 | 88 | 80 | 90 | 53 | 80 | 78.0346 | 75.3507 | 78 | 76 | 77 | 27.1 | 82 | 82.4 | [7 empty] | 135
| 1 | 6 | 80 | 80 | 84 | 80 | 90 | 52 | 70 | 82.5702 | 72.9017 | 77 | 74 | 78 | 27.1 | 73 | 74.6
| Mean |  | 80.3333 | 75 | 84.5 | 80 | 87.5 | 52.6667 | 73.3333 | 80.469 | 75.5799 | 76.8333 | 75.1667 | 78 | 27.1 | 78.6667 | 77.7833
| Median |  | 80.5 | 80 | 84 | 80 | 87.5 | 53 | 70 | 81.0131 | 75.5166 | 77 | 76 | 78 | 27.1 | 81.5 | 76.15
| Std Dev. |  | 1.3663 | 8.3666 | 1.8708 | 0 | 2.7386 | 0.5164 | 5.164 | 1.8791 | 2.812 | 1.472 | 1.3292 | 0.89443 | 0.063246 | 5.2409 | 3.6493
| Rel.Std.Dev. |  | 1.701% | 11.16% | 2.214% | 0% | 3.13% | 0.9805% | 7.042% | 2.335% | 3.721% | 1.916% | 1.768% | 1.147% | 0.2334% | 6.662% | 4.692%
| PDM3 |  | 2.699% | -4.119% | 8.026% | 2.273% | 11.86% | -32.67% | -6.25% | 2.872% | -3.378% | -1.775% | -3.906% | -0.284% | -65.36% | 0.5683% | -0.561%
| Z-Score (Absolute) |  | 0.66 | 0.72 | 1.73 | 0.57 | 2.5 | 6.47 | 1.15 | 0.69 | 0.57 | 0.24 | 0.67 | 0.06 | 13.06 | 0.23 | 0
| Analytical results for Ta in OREAS 550b (Certified Value 1.53 ppm) | [23 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 14 | 15 | 17 | 18 | 22 | 25 | 26 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [9 empty] | 2
2
| 1 | 1 | 1.5 | 1.7 | 1.9 | 3.2 | 1.5719 | 1.7 | < 10 | 1.483 | 1.4 | 1.4 | 1.5 | 1.06 | 0.8 | [9 empty] | 1
| 1 | 2 | 1.4 | 1.9 | 1.8 | 2.9 | 1.8925 | 1.5 | < 10 | 1.6585 | 1.4 | 1.6 | 1.6 | 1.04 | 1 | [9 empty] | 7
| 1 | 3 | 1.6 | 1.7 | 1.9 | 2.6 | 1.8681 | 1.4 | < 10 | 1.5005 | 1.4 | 1.6 | 1.6 | 0.99 | 0.8 | [9 empty] | 16
| 1 | 4 | 1.6 | 2.2 | 1.9 | 2.5 | 1.8401 | 1.4 | < 10 | 1.5427 | 1.5 | 1.3 | 1.5 | 1.04 | 0.7 | [9 empty] | 1.5326
| 1 | 5 | 1.6 | 1.7 | 1.8 | 2.4 | 1.6409 | 1.5 | < 10 | 1.4865 | 1.5 | 1.3 | 1.5 | 1.02 | 0.7 | [9 empty] | 136
| 1 | 6 | 1.5 | 1.8 | 1.9 | 2.4 | 1.7449 | 1.2 | < 10 | 1.3478 | 1.4 | 1.4 | 1.5 | 1.17 | 0.6
| Mean |  | 1.5333 | 1.8333 | 1.8667 | 2.6667 | 1.7597 | 1.45 |  | 1.5032 | 1.4333 | 1.4333 | 1.5333 | 1.0533 | 0.76667
| Median |  | 1.55 | 1.75 | 1.9 | 2.55 | 1.7925 | 1.45 |  | 1.4935 | 1.4 | 1.4 | 1.5 | 1.04 | 0.75
| Std Dev. |  | 0.08165 | 0.19664 | 0.05164 | 0.32042 | 0.13074 | 0.16432 |  | 0.10054 | 0.05164 | 0.13663 | 0.05164 | 0.06186 | 0.13663
| Rel.Std.Dev. |  | 5.325% | 10.73% | 2.766% | 12.02% | 7.429% | 11.33% |  | 6.689% | 3.603% | 9.532% | 3.368% | 5.873% | 17.82%
| PDM3 |  | 0.04636% | 19.62% | 21.8% | 73.99% | 14.82% | -5.391% |  | -1.922% | -6.478% | -6.478% | 0.04636% | -31.27% | -49.98%
| Z-Score (Absolute) |  | 0 | 0.89 | 0.99 | 3.38 | 0.67 | 0.25 | 10.33 | 0.09 | 0.3 | 0.3 | 0 | 1.43 | 2.28
| Analytical results for Tb in OREAS 550b (Certified Value 0.96 ppm) | [23 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 0.9 | 0.93 | 1.06 | 0.93739 | 1.03 | 0.9 | 0.99 | 1.1 | [14 empty] | 1
| 1 | 2 | 0.9 | 0.92 | 1.06 | 0.98369 | 1 | 0.9 | 0.92 | 1.12 | [14 empty] | 8
| 1 | 3 | 0.9 | 0.93 | 1.02 | 0.93075 | 1.02 | 0.9 | 0.95 | 1.06 | [14 empty] | 16
| 1 | 4 | 1 | 0.92 | 1.17 | 0.96135 | 0.95 | 0.9 | 0.97 | 1.11 | [14 empty] | 0.96125
| 1 | 5 | 0.9 | 0.94 | 1.07 | 0.92106 | 1.01 | 0.9 | 0.99 | 1.13 | [14 empty] | 137
| 1 | 6 | 1 | 0.9 | 1.15 | 0.90624 | 0.98 | 0.9 | 0.95 | 1.06
| Mean |  | 0.93333 | 0.92333 | 1.0883 | 0.94008 | 0.99833 | 0.9 | 0.96167 | 1.0967
| Median |  | 0.9 | 0.925 | 1.065 | 0.93407 | 1.005 | 0.9 | 0.96 | 1.105
| Std Dev. |  | 0.05164 | 0.013663 | 0.058452 | 0.028125 | 0.029269 | 0 | 0.027142 | 0.030111
| Rel.Std.Dev. |  | 5.533% | 1.48% | 5.371% | 2.992% | 2.932% | 0% | 2.822% | 2.746%
| PDM3 |  | -2.904% | -3.944% | 13.22% | -2.202% | 3.858% | -6.372% | 0.04339% | 14.09%
| Z-Score (Absolute) |  | 0.32 | 0.5 | 2.47 | 0.19 | 0.85 | 0.91 | 0.19 | 2.62
| Analytical results for Te in OREAS 550b (Indicative Value < 1 ppm) | [23 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 03 | 04 | 06 | 09 | 25 | 26 | [16 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [16 empty] | 2
2
| 1 | 1 | < 6 | < 1 | 2 | < 1 | < 0.1 | < 1 | [16 empty] | 1
| 1 | 2 | < 6 | < 1 | 1 | < 1 | < 0.1 | < 1 | [16 empty] | 9
| 1 | 3 | < 6 | < 1 | < 1 | < 1 | < 0.1 | < 1 | [16 empty] | 16
| 1 | 4 | < 6 | < 1 | < 1 | < 1 | < 0.1 | < 1 | [16 empty] | < 1
| 1 | 5 | < 6 | < 1 | < 1 | < 1 | < 0.1 | < 1 | [16 empty] | 26
| 1 | 6 | < 6 | < 1 | < 1 | < 1 | < 0.1 | < 1
| Mean |  |  |  | 1.5
| Median |  |  |  | 1.5
| Std Dev. |  |  |  | 0.70711
| Rel.Std.Dev. |  |  |  | 47.14%
PDM3
| Z-Score (Absolute) |  | 10.11 | 0 | 1.35 | 0 | 1.82 | 0
| Analytical results for Th in OREAS 550b (Certified Value 13.7 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 1
2
| 1 | 1 | 13.6 | 13.2 | 14.1 | 16.7 | 13.7948 | 13 | 12.1802 | 13.7 | 13.5 | 13.5 | 14.1 | 13.6 | [10 empty] | 1
| 1 | 2 | 14.1 | 13.6 | 13.9 | 15.5 | 13.7752 | 13.2 | 12.6532 | 14 | 13.7 | 13.4 | 14.2 | 14.4 | [10 empty] | 10
| 1 | 3 | 13.6 | 13 | 14.1 | 15.6 | 14.1063 | 13.6 | 12.419 | 13.8 | 13.7 | 13.6 | 12.6 | 14 | [10 empty] | 16
| 1 | 4 | 13.4 | 13.5 | 14.5 | 15.9 | 14.0744 | 13 | 12.26 | 14 | 13.5 | 13.3 | 13.7 | 14.3 | [10 empty] | 13.6757
| 1 | 5 | 14.1 | 13.9 | 13.8 | 15.4 | 13.5892 | 13 | 12.5236 | 13.9 | 13.7 | 13.1 | 13.4 | 14.1 | [10 empty] | 138
| 1 | 6 | 13.7 | 13.8 | 13.9 | 15.5 | 13.8006 | 12.7 | 12.5124 | 14.1 | 13.5 | 13.5 | 13.4 | 13.8
| Mean |  | 13.75 | 13.5 | 14.05 | 15.7667 | 13.8568 | 13.0833 | 12.4247 | 13.9167 | 13.6 | 13.4 | 13.5667 | 14.0333
| Median |  | 13.65 | 13.55 | 14 | 15.55 | 13.7977 | 13 | 12.4657 | 13.95 | 13.6 | 13.45 | 13.55 | 14.05
| Std Dev. |  | 0.2881 | 0.34641 | 0.251 | 0.48854 | 0.19742 | 0.29944 | 0.17699 | 0.1472 | 0.10954 | 0.17889 | 0.58195 | 0.30111
| Rel.Std.Dev. |  | 2.095% | 2.566% | 1.786% | 3.099% | 1.425% | 2.289% | 1.424% | 1.058% | 0.8055% | 1.335% | 4.29% | 2.146%
| PDM3 |  | 0.5435% | -1.285% | 2.737% | 15.29% | 1.324% | -4.331% | -9.147% | 1.762% | -0.5534% | -2.016% | -0.7971% | 2.615%
| Z-Score (Absolute) |  | 0.2 | 0.46 | 0.98 | 5.46 | 0.47 | 1.54 | 3.26 | 0.63 | 0.2 | 0.72 | 0.28 | 0.94
| Analytical results for Ti in OREAS 550b (Certified Value 0.374 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.39 | 0.378 | 0.38 | 0.37 | 0.36 | 0.38 | 0.36 | 0.35174 | 0.378 | 0.366 | 0.372 | 0.366 | 0.372 | 0.37198 | 0.36 | 0.38 | 0.3651 | 0.36 | 0.12 | 0.37 | 0.38 |  | 1
| 1 | 2 | 0.4 | 0.384 | 0.38 | 0.38 | 0.35 | 0.39 | 0.37 | 0.36972 | 0.372 | 0.384 | 0.378 | 0.384 | 0.372 | 0.39387 | 0.37 | 0.38 | 0.3693 | 0.36 | 0.12 | 0.37 | 0.39 |  | 25
| 1 | 3 | 0.39 | 0.372 | 0.38 | 0.37 | 0.36 | 0.38 | 0.35 | 0.38277 | 0.408 | 0.378 | 0.378 | 0.396 | 0.366 | 0.38194 | 0.37 | 0.37 | 0.3704 | 0.36 | 0.12 | 0.37 | 0.38 |  | 16
| 1 | 4 | 0.43 | 0.384 | 0.38 | 0.37 | 0.38 | 0.39 | 0.36 | 0.3735 | 0.408 | 0.372 | 0.372 | 0.366 | 0.378 | 0.37911 | 0.37 | 0.37 | 0.3708 | 0.35 | 0.12 | 0.35 | 0.37 |  | 0.37423
| 1 | 5 | 0.39 | 0.384 | 0.38 | 0.38 | 0.37 | 0.38 | 0.37 | 0.34908 | 0.408 | 0.372 | 0.378 | 0.366 | 0.378 | 0.38037 | 0.37 | 0.4 | 0.3705 | 0.35 | 0.11 | 0.34 | 0.37 |  | 139
| 1 | 6 | 0.39 | 0.372 | 0.38 | 0.38 | 0.37 | 0.38 | 0.36 | 0.36425 | 0.384 | 0.378 | 0.378 | 0.372 | 0.372 | 0.37663 | 0.36 | 0.39 | 0.3702 | 0.36 | 0.12 | 0.35 | 0.39
| Mean |  | 0.39833 | 0.379 | 0.38 | 0.375 | 0.365 | 0.38333 | 0.36167 | 0.36518 | 0.393 | 0.375 | 0.376 | 0.375 | 0.373 | 0.38065 | 0.36667 | 0.38167 | 0.36938 | 0.35667 | 0.11833 | 0.35833 | 0.38
| Median |  | 0.39 | 0.381 | 0.38 | 0.375 | 0.365 | 0.38 | 0.36 | 0.36699 | 0.396 | 0.375 | 0.378 | 0.369 | 0.372 | 0.37974 | 0.37 | 0.38 | 0.3703 | 0.36 | 0.12 | 0.36 | 0.38
| Std Dev. |  | 0.016021 | 0.0058992 | 6.0809e-17 | 0.0054772 | 0.010488 | 0.005164 | 0.0075277 | 0.012962 | 0.016864 | 0.0062929 | 0.0030984 | 0.012442 | 0.0045166 | 0.007353 | 0.005164 | 0.01169 | 0.0021591 | 0.005164 | 0.0040825 | 0.013292 | 0.0089443
| Rel.Std.Dev. |  | 4.022% | 1.557% | 1.6e-14% | 1.461% | 2.873% | 1.347% | 2.081% | 3.55% | 4.291% | 1.678% | 0.824% | 3.318% | 1.211% | 1.932% | 1.408% | 3.063% | 0.5845% | 1.448% | 3.45% | 3.709% | 2.354%
| PDM3 |  | 6.441% | 1.275% | 1.542% | 0.2059% | -2.466% | 2.433% | -3.357% | -2.419% | 5.016% | 0.2059% | 0.4731% | 0.2059% | -0.3286% | 1.716% | -2.021% | 1.987% | -1.295% | -4.693% | -68.38% | -4.248% | 1.542%
| Z-Score (Absolute) |  | 2.28 | 0.39 | 0.51 | 0.03 | 0.93 | 0.84 | 1.26 | 0.92 | 1.74 | 0 | 0.1 | 0 | 0.19 | 0.58 | 0.77 | 0.67 | 0.51 | 1.74 | 24.73 | 1.58 | 0.51
| Analytical results for Tl in OREAS 550b (Certified Value 0.6 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 17 | 18 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | 0.6 | 0.6 | 0.7 | 0.7 | 0.54393 | 0.6 | 0.5 | < 1 | 0.6 | 0.6 | [12 empty] | 1
| 1 | 2 | 0.7 | 0.5 | 0.7 | 0.7 | 0.58547 | 0.6 | 0.6 | < 1 | 0.5 | 0.6 | [12 empty] | 26
| 1 | 3 | 0.7 | 0.5 | 0.6 | 0.7 | 0.58845 | 0.6 | 0.6 | < 1 | < 0.5 | 0.6 | [12 empty] | 16
| 1 | 4 | 0.6 | 0.5 | 0.6 | 0.7 | 0.56539 | 0.6 | 0.6 | < 1 | 0.5 | 0.6 | [12 empty] | 0.59784
| 1 | 5 | 0.7 | 0.5 | 0.6 | 0.6 | 0.52992 | 0.7 | 0.5 | < 1 | 0.5 | 0.7 | [12 empty] | 140
| 1 | 6 | 0.6 | 0.6 | 0.6 | 0.7 | 0.54993 | 0.6 | 0.6 | < 1 | 0.5 | 0.6
| Mean |  | 0.65 | 0.53333 | 0.63333 | 0.68333 | 0.56052 | 0.61667 | 0.56667 |  | 0.52 | 0.61667
| Median |  | 0.65 | 0.5 | 0.6 | 0.7 | 0.55766 | 0.6 | 0.6 |  | 0.5 | 0.6
| Std Dev. |  | 0.054772 | 0.05164 | 0.05164 | 0.040825 | 0.023452 | 0.040825 | 0.05164 |  | 0.044721 | 0.040825
| Rel.Std.Dev. |  | 8.427% | 9.682% | 8.154% | 5.974% | 4.184% | 6.62% | 9.113% |  | 8.6% | 6.62%
| PDM3 |  | 8.726% | -10.79% | 5.938% | 14.3% | -6.242% | 3.15% | -5.214% |  | -13.02% | 3.15%
| Z-Score (Absolute) |  | 0.79 | 0.79 | 0.56 | 1.24 | 0.42 | 0.34 | 0.34 | 1.24 | 1.57 | 0.34
| Analytical results for Tm in OREAS 550b (Certified Value 0.49 ppm) | [23 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 0.5 | 0.46 | 0.46 | 0.45702 | 0.51 | < 1 | 0.48 | 0.5 | [14 empty] | 1
| 1 | 2 | 0.5 | 0.48 | 0.5 | 0.44785 | 0.51 | < 1 | 0.48 | 0.53 | [14 empty] | 12
| 1 | 3 | 0.4 | 0.48 | 0.51 | 0.41485 | 0.5 | < 1 | 0.46 | 0.47 | [14 empty] | 16
| 1 | 4 | 0.5 | 0.48 | 0.53 | 0.44884 | 0.49 | < 1 | 0.47 | 0.53 | [14 empty] | 0.48575
| 1 | 5 | 0.5 | 0.48 | 0.52 | 0.45684 | 0.51 | < 1 | 0.5 | 0.54 | [14 empty] | 141
| 1 | 6 | 0.5 | 0.48 | 0.51 | 0.46568 | 0.49 | < 1 | 0.43 | 0.48
| Mean |  | 0.48333 | 0.47667 | 0.505 | 0.44851 | 0.50167 |  | 0.47 | 0.50833
| Median |  | 0.5 | 0.48 | 0.51 | 0.45284 | 0.505 |  | 0.475 | 0.515
| Std Dev. |  | 0.040825 | 0.008165 | 0.02429 | 0.017721 | 0.0098319 |  | 0.023664 | 0.029269
| Rel.Std.Dev. |  | 8.447% | 1.713% | 4.81% | 3.951% | 1.96% |  | 5.035% | 5.758%
| PDM3 |  | -0.4974% | -1.87% | 3.963% | -7.666% | 3.277% |  | -3.242% | 4.649%
| Z-Score (Absolute) |  | 0.4 | 0.71 | 0.63 | 2.05 | 0.48 | 0.4 | 1.03 | 0.79
| Analytical results for U in OREAS 550b (Certified Value 9.12 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 2
3
| 1 | 1 | 9.3 | 8.7 | 8.6 | 10 | 9.1377 | 8.2 | 8.8367 | 9.55 | 9.3 | 8.9 | 10.4 | 9 | [10 empty] | 1
| 1 | 2 | 9.5 | 8.6 | 8.5 | 10 | 9.2348 | 8.2 | 8.621 | 9.75 | 9.6 | 8.9 | 10.3 | 10 | [10 empty] | 13
| 1 | 3 | 9.3 | 8.6 | 8.5 | 9 | 9.3816 | 8.4 | 8.7917 | 9.65 | 9.1 | 9 | 8.91 | 10 | [10 empty] | 16
| 1 | 4 | 9.4 | 8.8 | 8.5 | 10 | 9.276 | 8.1 | 8.7635 | 9.58 | 9.2 | 8.7 | 9.49 | 10 | [10 empty] | 9.117
| 1 | 5 | 9.4 | 9 | 8.6 | 9 | 9.1512 | 8 | 8.6665 | 9.71 | 9.2 | 8.6 | 9.06 | 10 | [10 empty] | 142
| 1 | 6 | 9.3 | 8.9 | 8.6 | 10 | 9.3087 | 8.2 | 8.5767 | 9.58 | 8.9 | 8.7 | 8.4 | 10
| Mean |  | 9.3667 | 8.7667 | 8.55 | 9.6667 | 9.2483 | 8.1833 | 8.7093 | 9.6367 | 9.2167 | 8.8 | 9.4267 | 9.8333
| Median |  | 9.35 | 8.75 | 8.55 | 10 | 9.2554 | 8.2 | 8.715 | 9.615 | 9.2 | 8.8 | 9.275 | 10
| Std Dev. |  | 0.08165 | 0.1633 | 0.054772 | 0.5164 | 0.093839 | 0.13292 | 0.10314 | 0.080416 | 0.23166 | 0.15492 | 0.79618 | 0.40825
| Rel.Std.Dev. |  | 0.8717% | 1.863% | 0.6406% | 5.342% | 1.015% | 1.624% | 1.184% | 0.8345% | 2.513% | 1.76% | 8.446% | 4.152%
| PDM3 |  | 2.738% | -3.843% | -6.219% | 6.029% | 1.44% | -10.24% | -4.472% | 5.7% | 1.093% | -3.477% | 3.396% | 7.857%
| Z-Score (Absolute) |  | 0.21 | 0.72 | 1.06 | 0.68 | 0.02 | 1.63 | 0.81 | 0.63 | 0.02 | 0.67 | 0.3 | 0.93
| Analytical results for V in OREAS 550b (Certified Value 214 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 210 | 210 | 234 | 210 | 204.1585 | 195 | 217 | 200 | 211.3212 | 216.7981 | 216 | 210 | 42.2 | 219 | 228 | [7 empty] | 1
| 1 | 2 | 215 | 209 | 227 | 210 | 204.079 | 194 | 215 | 190 | 219.6306 | 227.0506 | 217 | 210 | 42.5 | 204 | 225 | [7 empty] | 28
| 1 | 3 | 213 | 209 | 228 | 210 | 202.382 | 200 | 215 | 210 | 219.7025 | 216.3713 | 216 | 207 | 42.3 | 219 | 227 | [7 empty] | 16
| 1 | 4 | 221 | 213 | 232 | 210 | 200.583 | 210 | 217 | 200 | 216.6797 | 214.654 | 217 | 210 | 42.4 | 236 | 224 | [7 empty] | 213.6536
| 1 | 5 | 211 | 212 | 234 | 210 | 204.814 | 184 | 216 | 190 | 214.3308 | 215.6934 | 216 | 215 | 42.4 | 243 | 222 | [7 empty] | 143
| 1 | 6 | 217 | 207 | 232 | 210 | 204.2675 | 210 | 217 | 200 | 209.0018 | 213.3862 | 214 | 214 | 42.4 | 215 | 234
| Mean |  | 214.5 | 210 | 231.1667 | 210 | 203.3807 | 198.8333 | 216.1667 | 198.3333 | 215.1111 | 217.3256 | 216 | 211 | 42.3667 | 222.6667 | 226.6667
| Median |  | 214 | 209.5 | 232 | 210 | 204.1188 | 197.5 | 216.5 | 200 | 215.5053 | 216.0324 | 216 | 210 | 42.4 | 219 | 226
| Std Dev. |  | 4.0866 | 2.1909 | 2.9944 | 0 | 1.5974 | 10.0879 | 0.98319 | 7.5277 | 4.3878 | 4.9211 | 1.0954 | 2.9665 | 0.10328 | 14.3201 | 4.1793
| Rel.Std.Dev. |  | 1.905% | 1.043% | 1.295% | 0% | 0.7854% | 5.074% | 0.4548% | 3.795% | 2.04% | 2.264% | 0.5072% | 1.406% | 0.2438% | 6.431% | 1.844%
| PDM3 |  | 0.3961% | -1.71% | 8.197% | -1.71% | -4.808% | -6.937% | 1.176% | -7.171% | 0.6822% | 1.719% | 1.098% | -1.242% | -80.17% | 4.219% | 6.091%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.5 | 0.67 | 1.67 | 2.35 | 0.25 | 2.42 | 0.09 | 0.42 | 0.22 | 0.52 | 25.79 | 1.22 | 1.82
| Analytical results for W in OREAS 550b (Certified Value 3.4 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [10 empty] | 2
2
| 1 | 1 | 4 | < 100 | 3.1 | 4 | < 50 | 3.0689 | 2.6937 | 3 | 4 | < 5 | 3 | 3 | [10 empty] | 1
| 1 | 2 | 3 | < 100 | 3.8 | 4 | < 50 | 2.991 | 2.5872 | 4 | 3 | < 5 | 3 | 3 | [10 empty] | 14
| 1 | 3 | 4 | < 100 | 4.2 | 4 | < 50 | 3.272 | 2.5575 | 3 | 4 | < 5 | 4 | 3 | [10 empty] | 16
| 1 | 4 | 4 | < 100 | 6 | 4 | < 50 | 3.1828 | 2.7646 | 4 | 3 | < 5 | 3 | 3 | [10 empty] | 3.4021
| 1 | 5 | 5 | < 100 | 3.4 | 4 | < 50 | 2.9963 | 2.905 | 4 | 3 | < 5 | 4 | 3 | [10 empty] | 144
| 1 | 6 | 3 | < 100 | 3.4 | 4 | < 50 | 3.2249 | 2.9894 | 4 | 3 | < 5 | 3 | 7
| Mean |  | 3.8333 |  | 3.9833 | 4 |  | 3.1227 | 2.7496 | 3.6667 | 3.3333 |  | 3.3333 | 3.6667
| Median |  | 4 |  | 3.6 | 4 |  | 3.1258 | 2.7291 | 4 | 3 |  | 3 | 3
| Std Dev. |  | 0.75277 |  | 1.0591 | 0 |  | 0.12047 | 0.17222 | 0.5164 | 0.5164 |  | 0.5164 | 1.633
| Rel.Std.Dev. |  | 19.64% |  | 26.59% | 0% |  | 3.858% | 6.263% | 14.08% | 15.49% |  | 15.49% | 44.54%
| PDM3 |  | 12.68% |  | 17.08% | 17.57% |  | -8.214% | -19.18% | 7.777% | -2.021% |  | -2.021% | 7.777%
| Z-Score (Absolute) |  | 0.34 | 93.73 | 0.64 | 0.67 | 43.16 | 1.1 | 1.86 | 0 | 0.67 | 2.36 | 0.67 | 0
| Analytical results for Y in OREAS 550b (Certified Value 29.4 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [7 empty] | 1
2
| 1 | 1 | 30.3 | 34 | 31.1 | 27.3 | 30.4 | 30 | 28 | 29.1056 | 28.6604 | 30.9 | 29.7 | 30 | 24.2 | 18.6 | 29.4 | [7 empty] | 1
| 1 | 2 | 30.5 | 29 | 29.3 | 28.9 | 29.5 | 30 | 27 | 30.7074 | 29.5617 | 29.5 | 29.9 | 30 | 24.4 | 15.7 | 32.7 | [7 empty] | 15
| 1 | 3 | 30.2 | 27 | 24.8 | 29.3 | 28.6 | 30 | 28 | 30.6625 | 29.6331 | 29.1 | 29.5 | 31 | 24.8 | 20 | 29.4 | [7 empty] | 16
| 1 | 4 | 29.1 | 36 | 30.6 | 28.1 | 32.4 | 30 | 29 | 28.745 | 29.7769 | 28.6 | 29.3 | 31 | 24.1 | 18.6 | 30 | [7 empty] | 29.4088
| 1 | 5 | 29.1 | 28 | 29 | 28.9 | 29.9 | 30 | 29 | 28.5557 | 29.3126 | 28.6 | 29.3 | 30 | 23.4 | 18.6 | 31.1 | [7 empty] | 145
| 1 | 6 | 30 | 28 | 29.3 | 27.3 | 30.8 | 30 | 27 | 30.001 | 28.825 | 28.5 | 30.4 | 30 | 24.4 | 16.1 | 29.5
| Mean |  | 29.8667 | 30.3333 | 29.0167 | 28.3 | 30.2667 | 30 | 28 | 29.6295 | 29.2949 | 29.2 | 29.6833 | 30.3333 | 24.2167 | 17.9333 | 30.35
| Median |  | 30.1 | 28.5 | 29.3 | 28.5 | 30.15 | 30 | 28 | 29.5533 | 29.4371 | 28.85 | 29.6 | 30 | 24.3 | 18.6 | 29.75
| Std Dev. |  | 0.61536 | 3.7238 | 2.2266 | 0.86718 | 1.2925 | 0 | 0.89443 | 0.95673 | 0.45643 | 0.91652 | 0.42151 | 0.5164 | 0.46655 | 1.6705 | 1.3217
| Rel.Std.Dev. |  | 2.06% | 12.28% | 7.673% | 3.064% | 4.271% | 0% | 3.194% | 3.229% | 1.558% | 3.139% | 1.42% | 1.702% | 1.927% | 9.315% | 4.355%
| PDM3 |  | 1.557% | 3.144% | -1.333% | -3.77% | 2.917% | 2.01% | -4.79% | 0.7505% | -0.3872% | -0.71% | 0.9335% | 3.144% | -17.66% | -39.02% | 3.2%
| Z-Score (Absolute) |  | 0.25 | 0.74 | 0.65 | 1.41 | 0.67 | 0.39 | 1.72 | 0 | 0.35 | 0.45 | 0.06 | 0.74 | 5.73 | 12.38 | 0.76
| Analytical results for Yb in OREAS 550b (Certified Value 3.13 ppm) | [23 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 2.8 | 3.1 | 3.1 | 3.0491 | 3.4 | 3 | 2.6 | 3.3 | 2.9 | [13 empty] | 1
| 1 | 2 | 3.2 | 3 | 3.1 | 3.106 | 3.3 | 3 | 2.6 | 2.9 | 3.3 | [13 empty] | 16
| 1 | 3 | 2.2 | 3.1 | 3.1 | 2.9468 | 3.2 | 3 | 2.5 | 3.3 | 3.2 | [13 empty] | 16
| 1 | 4 | 3.2 | 3.1 | 3.5 | 2.8286 | 3.1 | 3 | 2.6 | 3.5 | 3.5 | [13 empty] | 3.1314
| 1 | 5 | 2.8 | 3.2 | 3.2 | 2.8608 | 3.2 | 3 | 2.7 | 3.3 | 3.3 | [13 empty] | 146
| 1 | 6 | 3.3 | 3.2 | 3.3 | 2.7569 | 3.1 | 3 | 2.6 | 3.4 | 3.2
| Mean |  | 2.9167 | 3.1167 | 3.2167 | 2.9247 | 3.2167 | 3 | 2.6 | 3.2833 | 3.2333
| Median |  | 3 | 3.1 | 3.15 | 2.9038 | 3.2 | 3 | 2.6 | 3.3 | 3.25
| Std Dev. |  | 0.41191 | 0.075277 | 0.16021 | 0.13439 | 0.1169 | 0 | 0.063246 | 0.20412 | 0.19664
| Rel.Std.Dev. |  | 14.12% | 2.415% | 4.981% | 4.595% | 3.634% | 0% | 2.433% | 6.217% | 6.082%
| PDM3 |  | -6.858% | -0.4711% | 2.722% | -6.602% | 2.722% | -4.197% | -16.97% | 4.851% | 3.255%
| Z-Score (Absolute) |  | 1.16 | 0 | 0.58 | 1.11 | 0.58 | 0.67 | 2.99 | 0.96 | 0.67
| Analytical results for Zn in OREAS 550b (Indicative Value 31.5 ppm) | [23 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS |  |  |  | 1
1
| 1 | 1 | 30.67 | 90 | 50 | 30 | 16 | 40 | < 20 | 20 | 80 | < 100 | 80 | < 20 | 23 | 28 | 43 | < 100 | 26.5 | 30 | 28 |  |  |  | 1
| 1 | 2 | 36.6 | < 50 | 30 | 30 | 16 | 30 | < 20 | 20 | 70 | < 100 | 60 | 50 | 25 | 27 | 40 | < 100 | 26.4 | 29 | 27 |  |  |  | 17
| 1 | 3 | 40 | < 50 | 30 | 30 | 16 | 30 | < 20 | 20 | 70 | < 100 | 60 | < 20 | 37 | 29 | 39 | < 100 | 26.4 | 44 | 29 |  |  |  | 16
| 1 | 4 | 32.35 | 50 | 50 | 30 | 16 | 30 | < 20 | < 20 | 60 | < 100 | 70 | < 20 | 26 | 26 | 47 | < 100 | 26.4 | 30 | 30 |  |  |  | 31.4668
| 1 | 5 | 34.98 | 60 | 30 | 30 | 16 | 40 | < 20 | 20 | 60 | < 100 | 60 | < 20 | 26 | 30 | 55 | < 100 | 26.4 | 36 | 29 |  |  |  | 27
| 1 | 6 | 29.51 | < 50 | 50 | 30 | 16 | 30 | < 20 | 20 | 60 | < 100 | 60 | < 20 | 26 | 29 | 40 | < 100 | 26.4 | 29 | 32
| Mean |  | 34.0183 | 66.6667 | 40 | 30 | 16 | 33.3333 |  | 20 | 66.6667 |  | 65 | 50 | 27.1667 | 28.1667 | 44 |  | 26.4167 | 33 | 29.1667
| Median |  | 33.665 | 60 | 40 | 30 | 16 | 30 |  | 20 | 65 |  | 60 | 50 | 26 | 28.5 | 41.5 |  | 26.4 | 30 | 29
| Std Dev. |  | 3.9418 | 20.8167 | 10.9545 | 0 | 0 | 5.164 |  | 0 | 8.165 |  | 8.3666 |  | 4.9565 | 1.472 | 6.1319 |  | 0.040825 | 6 | 1.7224
| Rel.Std.Dev. |  | 11.59% | 31.22% | 27.39% | 0% | 0% | 15.49% |  | 0% | 12.25% |  | 12.87% |  | 18.24% | 5.226% | 13.94% |  | 0.1545% | 18.18% | 5.905%
| PDM3 |  | 8.109% | 111.9% | 27.12% | -4.661% | -49.15% | 5.932% |  | -36.44% | 111.9% |  | 106.6% | 58.9% | -13.67% | -10.49% | 39.83% |  | -16.05% | 4.872% | -7.31%
| Z-Score (Absolute) |  | 0.06 | 0.79 | 0.43 | 0.18 | 1.04 | 0.02 | 1.41 | 0.9 | 2.06 | 1.04 | 1.96 | 1 | 0.36 | 0.3 | 0.67 | 1.04 | 0.4 | 0 | 0.23
| Analytical results for Zr in OREAS 550b (Indicative Value 174 ppm) | [23 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 15 | 17 | 19 | 22 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | [18 empty] | 0
0
| 1 | 1 | 172.7107 | 184 | 168 | 16.1 | [18 empty] | 1
| 1 | 2 | 174.0607 | 182 | 169 | 16.2 | [18 empty] | 22
| 1 | 3 | 173.1421 | 175 | 172 | 16.1 | [18 empty] | 16
| 1 | 4 | 174.5734 | 182 | 173 | 16.2 | [18 empty] | 174.1111
| 1 | 5 | 174.2613 | 174 | 168 | 16.2 | [18 empty] | 28
| 1 | 6 | 174.252 | 175 | 169 | 16.2
| Mean |  | 173.8334 | 178.6667 | 169.8333 | 16.1667
| Median |  | 174.1564 | 178.5 | 169 | 16.2
| Std Dev. |  | 0.73429 | 4.4572 | 2.137 | 0.05164
| Rel.Std.Dev. |  | 0.4224% | 2.495% | 1.258% | 0.3194%
| PDM3 |  | -0.1595% | 2.616% | -2.457% | -90.71%
| Z-Score (Absolute) |  | 0.31 | 1.04 | 0.31 | 23.77
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C649: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C667: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C703: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C721: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C739: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C757: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C775: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C812: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C830: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C848: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C866: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C884: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C902: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C920: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C938: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C956: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C974: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C992: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1010: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1100: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 550b (Certified Value 0.233 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.22 | 0.46 | 0.207 | 0.23 | < 0.5 | 0.23 | 0.27943 | 0.28 | 0.22 | 0.61 | 0.92 | 0.19 | 0.23055 | 0.2 | 0.24 | 0.19 | < 0.3 | 0.5 | 0.2 | 0.89 | < 1 | 0.21 | 0.22 | 0.5 |  |  | 1
| 1 | 2 | 0.44 | 0.216 | 0.236 | 0.22 | < 0.5 | 0.21 | 0.25977 | 0.22 | 0.29 | 0.2 | 0.19 | 0.4 | 0.22351 | 0.2 | 0.25 | 0.49 | 0.4 | 0.6 | 0.2 | 0.89 | < 1 | 0.19 | 0.22 | 1.9 |  |  | 4
| 1 | 3 | 0.45 | 0.232 | 0.302 | 0.14 | < 0.5 | 0.24 | 0.27384 | 0.24 | 0.23 | 0.38 | 0.4 | 0.26 | 0.25264 | 0.2 | 0.24 | 0.28 | < 0.3 | 0.6 | 0.2 | 0.88 | < 1 | 0.21 | 0.21 | < 0.5 |  |  | 16
| 1 | 4 | 0.23 | 0.187 | 0.219 | 0.25 | < 0.5 | 0.2 | 0.25458 | 0.44 | 0.21 | 0.22 | 0.24 | 0.21 | 0.24285 | 0.3 | 0.22 | 0.27 | 0.3 | 0.3 | 0.2 | 0.87 | < 1 | 0.21 | 0.22 | < 0.5 |  |  | 0.23312
| 1 | 5 | 0.21 | 0.206 | 0.279 | 0.24 | < 0.5 | 0.23 | 0.2696 | 0.2 | 0.3 | 0.22 | 0.21 | 0.26 | 0.22513 | 0.2 | 0.26 | 0.33 | < 0.3 | 0.5 | 0.3 | 0.86 | < 1 | 0.2 | 0.21 | < 0.5 |  |  | 148
| 1 | 6 | 0.28 | 0.18 | 0.154 | 0.24 | < 0.5 | 0.19 | 0.26007 | 0.22 | 0.22 | 0.23 | 0.21 | 0.69 | 0.22846 | 0.3 | 0.22 | 0.3 | < 0.3 | 0.2 | 0.2 | 0.87 | < 1 | 0.22 | 0.22 | 0.9
| Mean |  | 0.305 | 0.24683 | 0.23283 | 0.22 |  | 0.21667 | 0.26622 | 0.26667 | 0.245 | 0.31 | 0.36167 | 0.335 | 0.23386 | 0.23333 | 0.23833 | 0.31 | 0.35 | 0.45 | 0.21667 | 0.87667 |  | 0.20667 | 0.21667 | 1.1
| Median |  | 0.255 | 0.211 | 0.2275 | 0.235 |  | 0.22 | 0.26484 | 0.23 | 0.225 | 0.225 | 0.225 | 0.26 | 0.22951 | 0.2 | 0.24 | 0.29 | 0.35 | 0.5 | 0.2 | 0.875 |  | 0.21 | 0.22 | 0.9
| Std Dev. |  | 0.11113 | 0.10614 | 0.052913 | 0.040497 |  | 0.019664 | 0.0095818 | 0.089144 | 0.03937 | 0.161 | 0.28407 | 0.18876 | 0.011461 | 0.05164 | 0.016021 | 0.0998 | 0.070711 | 0.16432 | 0.040825 | 0.012111 |  | 0.010328 | 0.005164 | 0.72111
| Rel.Std.Dev. |  | 36.44% | 43% | 22.73% | 18.41% |  | 9.076% | 3.599% | 33.43% | 16.07% | 51.93% | 78.55% | 56.35% | 4.901% | 22.13% | 6.722% | 32.19% | 20.2% | 36.51% | 18.84% | 1.381% |  | 4.997% | 2.383% | 65.56%
| PDM3 |  | 30.83% | 5.882% | -0.1235% | -5.629% |  | -7.058% | 14.2% | 14.39% | 5.096% | 32.98% | 55.14% | 43.7% | 0.3154% | 0.09096% | 2.236% | 32.98% | 50.14% | 93.03% | -7.058% | 276.1% |  | -11.35% | -7.058% | 371.9%
| Z-Score (Absolute) |  | 1.11 | 0.06 | 0.35 | 0.61 | 0 | 0.67 | 0.33 | 0.34 | 0.1 | 1.21 | 2.26 | 1.72 | 0.33 | NA | 0.24 | 1.21 | NA | NA | NA | 12.68 | 5.06 | 0.88 | 0.67 | NA
Results from laboratories 16, 19, 20, 21 and 26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 550b (Certified Value 1.31 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 1.3878 | 1.25 | 1.74 | 1.89 | 1.576 | 1.44 | 1.1 | 1.3884 | 1.28 | 1.16 | 1.22 | 1.36 | 1.2 | 1.4574 | 1.14 | 1.31 | 2.1789 | 1.22 | 1.27 | 0.81 | 1.4149 | 1.339 | 2.13 | 1.264 |  |  | 1
| 1 | 2 | 1.4151 | 1.36 | 1.47 | 1.91 | 1.603 | 1.45 | 1.14 | 1.3251 | 1.28 | 1.24 | 1.2 | 1.39 | 1.16 | 1.404 | 1.17 | 1.32 | 2.1377 | 1.23 | 1.28 | 0.82 | 1.3803 | 1.302 | 1.98 | 1.265 |  |  | #N/A
| 1 | 3 | 1.3488 | 1.43 | 1.4 | 1.89 | 1.412 | 1.45 | 1.1 | 1.3651 | 1.25 | 1.25 | 1.24 | 1.34 | 1.18 | 1.4279 | 1.14 | 1.34 | 2.1113 | 1.22 | 1.28 | 0.83 | 1.3777 | 1.254 | 2.24 | 1.314 |  |  | 16
| 1 | 4 | 1.3409 | 1.42 | 1.66 | 1.89 | 1.371 | 1.44 | 1.15 | 1.3793 | 1.27 | 1.22 | 1.17 | 1.36 | 1.16 | 1.4387 | 1.17 | 1.35 | 2.1916 | 1.28 | 1.27 | 0.83 | 1.4153 | 1.371 | 1.96 | 1.27 |  |  | 1.3109
| 1 | 5 | 1.4113 | 1.22 | 1.39 | 1.88 | 1.541 | 1.44 | 1.13 | 1.3832 | 1.26 | 1.18 | 1.19 | 1.34 | 1.17 | 1.4169 | 1.15 | 1.31 | 2.0689 | 1.27 | 1.27 | 0.81 | 1.4176 | 1.254 | 1.8 | 1.253 |  |  | 149
| 1 | 6 | 1.3857 | 1.31 | 1.7 | 1.9 | 1.321 | 1.42 | 1.1 | 1.3357 | 1.28 | 1.22 | 1.18 | 1.32 | 1.16 | 1.4098 | 1.14 | 1.33 | 2.1178 | 1.35 | 1.25 | 0.82 | 1.3782 | 1.397 | 2.21 | 1.278
| Mean |  | 1.3816 | 1.3317 | 1.56 | 1.8933 | 1.4707 | 1.44 | 1.12 | 1.3628 | 1.27 | 1.2117 | 1.2 | 1.3517 | 1.1717 | 1.4258 | 1.1517 | 1.3267 | 2.1344 | 1.2617 | 1.27 | 0.82 | 1.3973 | 1.3195 | 2.0533 | 1.274
| Median |  | 1.3867 | 1.335 | 1.565 | 1.89 | 1.4765 | 1.44 | 1.115 | 1.3722 | 1.275 | 1.22 | 1.195 | 1.35 | 1.165 | 1.4224 | 1.145 | 1.325 | 2.1277 | 1.25 | 1.27 | 0.82 | 1.3976 | 1.3205 | 2.055 | 1.2675
| Std Dev. |  | 0.030958 | 0.087044 | 0.15786 | 0.010328 | 0.11775 | 0.010954 | 0.022804 | 0.026486 | 0.012649 | 0.034881 | 0.026077 | 0.024014 | 0.016021 | 0.019894 | 0.01472 | 0.01633 | 0.045524 | 0.050365 | 0.010954 | 0.0089443 | 0.020393 | 0.059876 | 0.16943 | 0.021232
| Rel.Std.Dev. |  | 2.241% | 6.536% | 10.12% | 0.5455% | 8.007% | 0.7607% | 2.036% | 1.943% | 0.996% | 2.879% | 2.173% | 1.777% | 1.367% | 1.395% | 1.278% | 1.231% | 2.133% | 3.992% | 0.8626% | 1.091% | 1.459% | 4.538% | 8.251% | 1.667%
| PDM3 |  | 5.392% | 1.583% | 19% | 44.43% | 12.19% | 9.846% | -14.56% | 3.958% | -3.122% | -7.571% | -8.461% | 3.108% | -10.62% | 8.761% | -12.15% | 1.201% | 62.81% | -3.757% | -3.122% | -37.45% | 6.592% | 0.6544% | 56.63% | -2.816%
| Z-Score (Absolute) |  | 0.34 | 0.02 | 1.5 | 3.67 | 0.92 | 0.72 | 1.36 | 0.22 | 0.38 | 0.76 | 0.84 | 0.15 | 1.02 | 0.63 | 1.15 | 0.02 | 5.23 | 0.44 | 0.38 | 3.31 | 0.44 | 0.06 | 4.71 | 0.36
| Analytical results for As in OREAS 550b (Certified Value 58 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 1
| 1 | 1 | 59 | 54.4 | 52.9 | 57 | 58.4 | 78 | 56 | 60 | 60.809 | 62.5 | 59.4 | 61.6 | 61 | 61.6 | 62.0169 | 54 | 57 | 61 | 57.3 | 55.9 | 57.2 | 38.3 | 52.586 | < 1 | 57.2 | 56 | 1
| 1 | 2 | 58 | 53.8 | 55.4 | 58 | 57.41 | 83 | 59 | 62 | 59.7191 | 63.5 | 59.9 | 63 | 61.5 | 58.9 | 59.3475 | 54.6 | 59 | 59 | 57.6 | 57.1 | 58.9 | 37.4 | 50.888 | < 1 | 57.2 | 60 | 5
| 1 | 3 | 59 | 55.3 | 54.6 | 58 | 57.76 | 72 | 59 | 59 | 61.0622 | 61.1 | 59.8 | 64.1 | 63.3 | 59.1 | 60.0021 | 54.7 | 57 | 58 | 57.6 | 55.3 | 58 | 38.6 | 52.46 | < 1 | 56.7 | 62 | 16
| 1 | 4 | 60 | 54.2 | 54.7 | 58 | 59.63 | 75 | 59 | 62 | 61.8097 | 63.4 | 60.1 | 62.1 | 61.3 | 62.2 | 59.8254 | 54.3 | 57 | 61 | 57.9 | 57.6 | 57.7 | 37.5 | 54.124 | < 1 | 55.8 | 58 | 58.4178
| 1 | 5 | 59 | 53.6 | 53.7 | 58 | 57.31 | 77 | 64 | 61 | 58.2694 | 61.2 | 57.9 | 63.8 | 63.1 | 62.1 | 59.0636 | 54.1 | 57 | 60 | 57.9 | 56.9 | 58.8 | 38.4 | 52.402 | < 1 | 55.1 | 60 | 150
| 1 | 6 | 58 | 56.9 | 53.4 | 58 | 58.57 | 67 | 61 | 60 | 57.3828 | 61.2 | 59.8 | 62.9 | 60.2 | 65 | 59.138 | 56.1 | 58 | 60 | 56.8 | 59.3 | 58.8 | 38.4 | 51.913 | < 1 | 53.9 | 63
| Mean |  | 58.8333 | 54.7 | 54.1167 | 57.8333 | 58.18 | 75.3333 | 59.6667 | 60.6667 | 59.842 | 62.15 | 59.4833 | 62.9167 | 61.7333 | 61.4833 | 59.8989 | 54.6333 | 57.5 | 59.8333 | 57.5167 | 57.0167 | 58.2333 | 38.1 | 52.3955 |  | 55.9833 | 59.8333
| Median |  | 59 | 54.3 | 54.15 | 58 | 58.08 | 76 | 59 | 60.5 | 60.2641 | 61.85 | 59.8 | 62.95 | 61.4 | 61.85 | 59.5865 | 54.45 | 57 | 60 | 57.6 | 57 | 58.4 | 38.35 | 52.431 |  | 56.25 | 60
| Std Dev. |  | 0.75277 | 1.2296 | 0.93684 | 0.40825 | 0.87512 | 5.465 | 2.6583 | 1.2111 | 1.7223 | 1.1327 | 0.8085 | 0.95795 | 1.2209 | 2.264 | 1.103 | 0.76855 | 0.83666 | 1.169 | 0.41673 | 1.3977 | 0.70616 | 0.51381 | 1.052 |  | 1.3136 | 2.5626
| Rel.Std.Dev. |  | 1.28% | 2.248% | 1.731% | 0.7059% | 1.504% | 7.254% | 4.455% | 1.996% | 2.878% | 1.823% | 1.359% | 1.523% | 1.978% | 3.682% | 1.841% | 1.407% | 1.455% | 1.954% | 0.7245% | 2.451% | 1.213% | 1.349% | 2.008% |  | 2.346% | 4.283%
| PDM3 |  | 0.7113% | -6.364% | -7.363% | -1.001% | -0.4071% | 28.96% | 2.138% | 3.85% | 2.438% | 6.389% | 1.824% | 7.701% | 5.676% | 5.248% | 2.535% | -6.478% | -1.571% | 2.423% | -1.543% | -2.398% | -0.3158% | -34.78% | -10.31% |  | -4.167% | 2.423%
| Z-Score (Absolute) |  | 0.11 | 1.42 | 1.63 | 0.26 | 0.13 | 6.21 | 0.42 | 0.79 | 0.48 | 1.34 | 0.35 | 1.62 | 1.18 | 1.09 | 0.5 | 1.44 | 0.38 | 0.48 | 0.38 | 0.56 | 0.11 | 7.55 | 2.27 | 21.44 | 0.94 | 0.48
| Analytical results for Au in OREAS 550b (Certified Value < 0.02 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 02 | 03 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 20 | 24 | 25 | [13 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [13 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [13 empty] | 3
| 1 | 1 | 0.001 | 0.0083 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.005 | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | 1
| 1 | 2 | 0.001 | 0.0129 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.003 | < 0.0005 | < 0.01 | 0.0007 | [13 empty] | #N/A
| 1 | 3 | 0.001 | 0.0036 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.004 | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | 16
| 1 | 4 | 0.001 | 0.0074 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.004 | < 0.0005 | < 0.01 | 0.0006 | [13 empty] | < 0.02
| 1 | 5 |  | 0.0046 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.004 | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | 151
| 1 | 6 | 0.001 | 0.0022 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.004 | < 0.0005 | < 0.01 | < 0.0005
| Mean |  | 0.001 | 0.0065 | [7 empty] | 0.004 |  |  | 0.00065
| Median |  | 0.001 | 0.006 | [7 empty] | 0.004 |  |  | 0.00065
| Std Dev. |  | 0 | 0.0038864 | [7 empty] | 0.00063246 |  |  | 7.0711e-05
| Rel.Std.Dev. |  | 0% | 59.79% | [7 empty] | 15.81% |  |  | 10.88%
PDM3
| Z-Score (Absolute) |  | 0.57 | 0.21 | 0.67 | 0.64 | 0.71 | 0.71 | 0.71 | 0.71 | 0.71 | 0.14 | 0.67 | 0 | 0.66
| Analytical results for B in OREAS 550b (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 03 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 20 | 25 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [13 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.1g | [13 empty] | 1
| 1 | 1 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.719 | < 20 | < 10 | [13 empty] | 1
| 1 | 2 | < 20 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.8682 | < 20 | < 10 | [13 empty] | 1
| 1 | 3 | < 20 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.3544 | < 20 | < 10 | [13 empty] | 16
| 1 | 4 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.4619 | < 20 | < 10 | [13 empty] | < 10
| 1 | 5 | < 20 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.6291 | < 20 | < 10 | [13 empty] | 152
| 1 | 6 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 5.0934 | < 20 | < 10
| Mean |  |  | 4.5 | [8 empty] | 4.6877
| Median |  |  | 4.5 | [8 empty] | 4.6741
| Std Dev. |  |  | 0.54772 | [8 empty] | 0.26968
| Rel.Std.Dev. |  |  | 12.17% | [8 empty] | 5.753%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 550b (Certified Value 103 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 0
| 1 | 1 | 109 | 98.4 | 19.3 | 135 | 131 | 93 | 92 | 130.8133 | 100 | 100 | 100 | 110 | 100 | 116.5312 | 84 | 99 | 164 | 92 | 107 | 89.7 | 99.454 | 152 | 165 | 85 |  |  | 1
| 1 | 2 | 108 | 103.6 | 53.9 | 133 | 139 | 94 | 90 | 124.5772 | 100 | 100 | 100 | 110 | 100 | 111.9896 | 83.3 | 99 | 166 | 97 | 100 | 88.2 | 98.308 | 138 | 156 | 87 |  |  | #N/A
| 1 | 3 | 108 | 109.1 | 35 | 136 | 118 | 94 | 87 | 131.6034 | 100 | 100 | 100 | 110 | 100 | 113.3667 | 85 | 101 | 164 | 95 | 98 | 86.3 | 97.399 | 104 | 174 | 88 |  |  | 16
| 1 | 4 | 111 | 109.4 | 37.1 | 137 | 114 | 94 | 91 | 127.8437 | 100 | 100 | 90 | 110 | 100 | 112.3373 | 84.7 | 100 | 164 | 98 | 99 | 86 | 101.384 | 124 | 174 | 86 |  |  | 103.4893
| 1 | 5 | 108 | 96.1 | 31.1 | 137 | 130 | 94 | 88 | 135.1025 | 100 | 100 | 90 | 110 | 100 | 109.2374 | 84.3 | 98 | 168 | 103 | 100 | 87.1 | 101.069 | 97 | 174 | 83 |  |  | 153
| 1 | 6 | 105 | 99.6 | 19.7 | 138 | 111 | 93 | 90 | 121.8202 | 100 | 100 | 100 | 110 | 100 | 109.582 | 84.3 | 99 | 162 | 103 | 99 | 88.7 | 98.938 | 155 | 174 | 92
| Mean |  | 108.1667 | 102.7 | 32.6833 | 136 | 123.8333 | 93.6667 | 89.6667 | 128.6267 | 100 | 100 | 96.6667 | 110 | 100 | 112.174 | 84.2667 | 99.3333 | 164.6667 | 98 | 100.5 | 87.6667 | 99.4253 | 128.3333 | 169.5 | 86.8333
| Median |  | 108 | 101.6 | 33.05 | 136.5 | 124 | 94 | 90 | 129.3285 | 100 | 100 | 100 | 110 | 100 | 112.1635 | 84.3 | 99 | 164 | 97.5 | 99.5 | 87.65 | 99.196 | 131 | 174 | 86.5
| Std Dev. |  | 1.9408 | 5.6271 | 12.8469 | 1.7889 | 11.089 | 0.5164 | 1.8619 | 4.8768 | 0 | 0 | 5.164 | 0 | 0 | 2.6762 | 0.58878 | 1.0328 | 2.0656 | 4.3818 | 3.2711 | 1.4459 | 1.5576 | 24.3201 | 7.5299 | 3.0605
| Rel.Std.Dev. |  | 1.794% | 5.479% | 39.31% | 1.315% | 8.955% | 0.5513% | 2.076% | 3.791% | 0% | 0% | 5.342% | 0% | 0% | 2.386% | 0.6987% | 1.04% | 1.254% | 4.471% | 3.255% | 1.649% | 1.567% | 18.95% | 4.442% | 3.525%
| PDM3 |  | 4.52% | -0.7627% | -68.42% | 31.41% | 19.66% | -9.491% | -13.36% | 24.29% | -3.372% | -3.372% | -6.593% | 6.291% | -3.372% | 8.392% | -18.57% | -4.016% | 59.11% | -5.304% | -2.889% | -15.29% | -3.927% | 24.01% | 63.78% | -16.09%
| Z-Score (Absolute) |  | 0.54 | 0.18 | 4.46 | 2.39 | 1.58 | 0.42 | 0.69 | 1.9 | 0 | 0 | 0.22 | 0.66 | 0 | 0.81 | 1.04 | 0.04 | 4.29 | 0.13 | 0.03 | 0.82 | 0.04 | 1.88 | 4.61 | 0.87
| Analytical results for Be in OREAS 550b (Certified Value 1.63 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 3
| 1 | 1 | 1.72 | 1.8 | 1.86 | 4.9444 | 1.5 | 1.7 | 1.6613 | 1.59 | 1.52 | 1.58 | 1.68 | 1.54 | 1.7246 | 1.5 | 1.7 | 2 | < 2 | 1.986 | 1.5 | 1.71 | 1.86 | [5 empty] | 1
| 1 | 2 | 1.64 | 1.6 | 1.83 | 4.9423 | 1.45 | 1.6 | 1.5659 | 1.61 | 1.56 | 1.56 | 1.71 | 1.49 | 1.6718 | 1.3 | 1.7 | 2 | < 2 | 1.644 | 1.46 | 1.45 | 1.74 | [5 empty] | 6
| 1 | 3 | 1.67 | 1.6 | 1.87 | 4.983 | 1.43 | 1.8 | 1.6473 | 1.6 | 1.57 | 1.56 | 1.68 | 1.51 | 1.6994 | 1.4 | 1.6 | 2 | < 2 | 1.709 | 1.55 | 1.53 | 1.85 | [5 empty] | 16
| 1 | 4 | 1.71 | 1.7 | 1.87 | 4.9444 | 1.33 | 1.6 | 1.6691 | 1.6 | 1.56 | 1.52 | 1.69 | 1.49 | 1.7015 | 1.4 | 1.7 | 1 | < 2 | 1.933 | 1.57 | 1.56 | 1.78 | [5 empty] | 1.6285
| 1 | 5 | 1.66 | 1.5 | 1.87 | 4.92 | 1.41 | 1.8 | 1.5848 | 1.59 | 1.52 | 1.52 | 1.68 | 1.51 | 1.6762 | 1.6 | 1.7 | 2 | < 2 | 1.942 | 1.58 | 1.52 | 1.72 | [5 empty] | 154
| 1 | 6 | 1.62 | 1.8 | 1.89 | 4.9643 | 1.48 | 1.7 | 1.5296 | 1.58 | 1.56 | 1.52 | 1.66 | 1.49 | 1.6766 | 1.6 | 1.8 | 1 | < 2 | 1.712 | 1.54 | 1.51 | 1.65
| Mean |  | 1.67 | 1.6667 | 1.865 | 4.9497 | 1.4333 | 1.7 | 1.6097 | 1.595 | 1.5483 | 1.5433 | 1.6833 | 1.505 | 1.6917 | 1.4667 | 1.7 | 1.6667 |  | 1.821 | 1.5333 | 1.5467 | 1.7667
| Median |  | 1.665 | 1.65 | 1.87 | 4.9444 | 1.44 | 1.7 | 1.616 | 1.595 | 1.56 | 1.54 | 1.68 | 1.5 | 1.688 | 1.45 | 1.7 | 2 |  | 1.8225 | 1.545 | 1.525 | 1.76
| Std Dev. |  | 0.038987 | 0.12111 | 0.019748 | 0.021518 | 0.060222 | 0.089443 | 0.057552 | 0.010488 | 0.022286 | 0.026583 | 0.01633 | 0.019748 | 0.020503 | 0.12111 | 0.063246 | 0.5164 |  | 0.14843 | 0.045461 | 0.087788 | 0.080416
| Rel.Std.Dev. |  | 2.335% | 7.266% | 1.059% | 0.4347% | 4.202% | 5.261% | 3.575% | 0.6576% | 1.439% | 1.722% | 0.9701% | 1.312% | 1.212% | 8.257% | 3.72% | 30.98% |  | 8.151% | 2.965% | 5.676% | 4.552%
| PDM3 |  | 2.548% | 2.344% | 14.52% | 203.9% | -11.98% | 4.39% | -1.157% | -2.057% | -4.923% | -5.23% | 3.367% | -7.584% | 3.88% | -9.938% | 4.39% | 2.344% |  | 11.82% | -5.844% | -5.025% | 8.484%
| Z-Score (Absolute) |  | 0.23 | 0.21 | 1.64 | 23.97 | 1.48 | 0.45 | 0.21 | 0.31 | 0.65 | 0.69 | 0.33 | 0.96 | 0.39 | 1.24 | 0.45 | NA | 4.62 | 1.32 | 0.76 | 0.66 | 0.93
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 550b (Certified Value 1.91 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 1.99 | 1.84 | 1.96 | 2.08 | 1.93 | 1.92 | 1.8777 | 1.93 | 1.77 | 1.82 | 1.82 | 2.01 | 2.1332 | 1.94 | 1.91 | 2.22 | 2.39 | 1.8 | 1.8 | < 5 | 1.96 | 1.84 | 2.01 | 1.87 |  |  | 1
| 1 | 2 | 1.97 | 1.88 | 1.82 | 2.09 | 1.92 | 1.96 | 1.9231 | 1.92 | 1.82 | 1.84 | 1.86 | 1.92 | 1.9879 | 1.94 | 1.93 | 2.12 | 2.37 | 1.9 | 1.8 | < 5 | 1.893 | 1.82 | 2.13 | 1.84 |  |  | 7
| 1 | 3 | 2 | 1.92 | 1.82 | 2.04 | 2 | 1.86 | 1.8723 | 1.92 | 1.84 | 1.89 | 1.8 | 1.86 | 2.0894 | 1.96 | 1.92 | 2.11 | 2.45 | 1.8 | 1.7 | < 5 | 1.881 | 1.85 | 2.04 | 1.84 |  |  | 16
| 1 | 4 | 1.96 | 1.93 | 1.91 | 2.05 | 1.95 | 1.95 | 1.9632 | 1.93 | 1.82 | 1.88 | 1.72 | 1.96 | 2.1171 | 1.95 | 1.93 | 2.14 | 2.61 | 1.8 | 1.7 | < 5 | 1.895 | 1.92 | 2.18 | 1.81 |  |  | 1.9093
| 1 | 5 | 1.96 | 1.85 | 1.77 | 2.03 | 2.08 | 1.86 | 1.9638 | 1.91 | 1.77 | 1.92 | 1.8 | 1.96 | 2.0314 | 1.92 | 1.86 | 2.1 | 2.5 | 1.9 | 1.7 | < 5 | 1.851 | 1.93 | 2.05 | 1.84 |  |  | 155
| 1 | 6 | 1.93 | 1.78 | 1.88 | 2.08 | 1.94 | 1.9 | 1.8558 | 1.89 | 1.82 | 1.9 | 1.8 | 1.96 | 2.0203 | 1.96 | 1.96 | 2.11 | 2.64 | 1.9 | 1.8 | < 5 | 1.84 | 1.9 | 2.06 | 1.87
| Mean |  | 1.9683 | 1.8667 | 1.86 | 2.0617 | 1.97 | 1.9083 | 1.9093 | 1.9167 | 1.8067 | 1.875 | 1.8 | 1.945 | 2.0632 | 1.945 | 1.9183 | 2.1333 | 2.4933 | 1.85 | 1.75 |  | 1.8867 | 1.8767 | 2.0783 | 1.845
| Median |  | 1.965 | 1.865 | 1.85 | 2.065 | 1.945 | 1.91 | 1.9004 | 1.92 | 1.82 | 1.885 | 1.8 | 1.96 | 2.0604 | 1.945 | 1.925 | 2.115 | 2.475 | 1.85 | 1.75 |  | 1.887 | 1.875 | 2.055 | 1.84
| Std Dev. |  | 0.024833 | 0.055737 | 0.06957 | 0.024833 | 0.060663 | 0.043089 | 0.047547 | 0.015055 | 0.029439 | 0.037815 | 0.045607 | 0.050498 | 0.058314 | 0.015166 | 0.033116 | 0.044572 | 0.11219 | 0.054772 | 0.054772 |  | 0.042354 | 0.045898 | 0.063692 | 0.022583
| Rel.Std.Dev. |  | 1.262% | 2.986% | 3.74% | 1.204% | 3.079% | 2.258% | 2.49% | 0.7855% | 1.629% | 2.017% | 2.534% | 2.596% | 2.826% | 0.7797% | 1.726% | 2.089% | 4.5% | 2.961% | 3.13% |  | 2.245% | 2.446% | 3.065% | 1.224%
| PDM3 |  | 3.09% | -2.234% | -2.584% | 7.979% | 3.178% | -0.05208% | -6.459e-06% | 0.3844% | -5.377% | -1.798% | -5.726% | 1.868% | 8.06% | 1.868% | 0.4717% | 11.73% | 30.59% | -3.107% | -8.345% |  | -1.187% | -1.711% | 8.852% | -3.369%
| Z-Score (Absolute) |  | 0.66 | 0.56 | 0.64 | 1.78 | 0.68 | 0.06 | 0.04 | 0.04 | 1.28 | 0.46 | 1.36 | 0.38 | 1.8 | 0.38 | 0.06 | 2.65 | 6.97 | 0.76 | 1.96 | 7.05 | 0.32 | 0.44 | 1.98 | 0.82
| Analytical results for Ca in OREAS 550b (Certified Value 1.32 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 25 | 26 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 3
| 1 | 1 | 1.23 | 1.18 | 1.35 | 1.23 | 1.3162 | 1.574 | 1.32 | 1.35 | 1.328 | 1.34 | 1.37 | 1.39 | 1.34 | 1.36 | 1.371 | 1.29 | 1.3 | 1.33 | 1.3391 | 1.27 | 1.15 | 1.32 | 1.387 | 1.21 | 1.361 |  | 1
| 1 | 2 | 1.25 | 1.17 | 1.25 | 1.26 | 1.3127 | 1.642 | 1.3 | 1.36 | 1.3056 | 1.38 | 1.39 | 1.36 | 1.36 | 1.32 | 1.3106 | 1.31 | 1.29 | 1.29 | 1.3541 | 1.27 | 1.17 | 1.33 | 1.401 | 1.26 | 1.369 |  | #N/A
| 1 | 3 | 1.19 | 1.19 | 1.28 | 1.27 | 1.3133 | 1.42 | 1.32 | 1.34 | 1.326 | 1.35 | 1.39 | 1.35 | 1.34 | 1.32 | 1.3509 | 1.29 | 1.3 | 1.29 | 1.3374 | 1.24 | 1.17 | 1.3 | 1.387 | 1.29 | 1.368 |  | 16
| 1 | 4 | 1.19 | 1.18 | 1.24 | 1.23 | 1.3161 | 1.374 | 1.32 | 1.36 | 1.3619 | 1.37 | 1.39 | 1.31 | 1.36 | 1.31 | 1.3667 | 1.3 | 1.29 | 1.33 | 1.346 | 1.29 | 1.17 | 1.33 | 1.422 | 1.28 | 1.356 |  | 1.3184
| 1 | 5 | 1.26 | 1.15 | 1.22 | 1.26 | 1.3156 | 1.57 | 1.3 | 1.35 | 1.3631 | 1.35 | 1.37 | 1.33 | 1.32 | 1.31 | 1.3326 | 1.3 | 1.29 | 1.33 | 1.3476 | 1.28 | 1.16 | 1.33 | 1.401 | 1.23 | 1.365 |  | 156
| 1 | 6 | 1.23 | 1.21 | 1.29 | 1.24 | 1.3191 | 1.358 | 1.29 | 1.33 | 1.2917 | 1.35 | 1.4 | 1.33 | 1.34 | 1.31 | 1.3424 | 1.29 | 1.29 | 1.32 | 1.3378 | 1.31 | 1.16 | 1.31 | 1.429 | 1.24 | 1.379
| Mean |  | 1.225 | 1.18 | 1.2717 | 1.2483 | 1.3155 | 1.4897 | 1.3083 | 1.3483 | 1.3294 | 1.3567 | 1.385 | 1.345 | 1.3433 | 1.3217 | 1.3457 | 1.2967 | 1.2933 | 1.315 | 1.3437 | 1.2767 | 1.1633 | 1.32 | 1.4045 | 1.2517 | 1.3663
| Median |  | 1.23 | 1.18 | 1.265 | 1.25 | 1.3158 | 1.495 | 1.31 | 1.35 | 1.327 | 1.35 | 1.39 | 1.34 | 1.34 | 1.315 | 1.3467 | 1.295 | 1.29 | 1.325 | 1.3426 | 1.275 | 1.165 | 1.325 | 1.401 | 1.25 | 1.3665
| Std Dev. |  | 0.029496 | 0.02 | 0.046224 | 0.017224 | 0.0023022 | 0.12028 | 0.013292 | 0.01169 | 0.028971 | 0.015055 | 0.012247 | 0.028107 | 0.015055 | 0.019408 | 0.02247 | 0.008165 | 0.005164 | 0.019748 | 0.006698 | 0.023381 | 0.008165 | 0.012649 | 0.01757 | 0.030605 | 0.0078401
| Rel.Std.Dev. |  | 2.408% | 1.695% | 3.635% | 1.38% | 0.175% | 8.074% | 1.016% | 0.867% | 2.179% | 1.11% | 0.8843% | 2.09% | 1.121% | 1.468% | 1.67% | 0.6297% | 0.3993% | 1.502% | 0.4985% | 1.831% | 0.7019% | 0.9583% | 1.251% | 2.445% | 0.5738%
| PDM3 |  | -7.082% | -10.49% | -3.542% | -5.312% | -0.2171% | 12.99% | -0.7607% | 2.273% | 0.8361% | 2.905% | 5.055% | 2.021% | 1.894% | 0.2507% | 2.074% | -1.646% | -1.898% | -0.255% | 1.919% | -3.163% | -11.76% | 0.1243% | 6.534% | -5.059% | 3.639%
| Z-Score (Absolute) |  | 2.26 | 3.33 | 1.15 | 1.71 | 0.11 | 4.04 | 0.28 | 0.67 | 0.22 | 0.87 | 1.55 | 0.59 | 0.56 | 0.04 | 0.61 | 0.56 | 0.63 | 0.12 | 0.56 | 1.03 | 3.73 | 0 | 2.01 | 1.63 | 1.1
| Analytical results for Cd in OREAS 550b (Certified Value 0.2 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.2 | 0.21 | 0.17 | 0.2 | < 1 | 0.25 | 0.22 | 0.2199 | 0.2 | 0.19 | 0.18 | 0.18 | 0.18 | 0.23535 | 0.19 | 0.19 |  | 0.2 | 0.2 | < 1 | 0.214 | 0.18 | 0.12 | 0.18 |  |  | 1
| 1 | 2 | 0.22 | 0.26 | 0.25 | 0.19 | < 1 | 0.22 | 0.2 | 0.18811 | 0.2 | 0.2 | 0.19 | 0.18 | 0.21 | 0.22746 | 0.21 | 0.19 | 0.23 | 0.2 | 0.2 | < 1 | 0.22 | 0.19 | 0.13 | 0.19 |  |  | 8
| 1 | 3 | 0.2 | 0.2 | 0.2 | 0.19 | < 1 | 0.24 | 0.21 | 0.17973 | 0.21 | 0.2 | 0.19 | 0.18 | 0.19 | 0.22524 | 0.22 | 0.19 | 0.2 | 0.2 | 0.2 | < 1 | 0.217 | 0.2 | 0.13 | 0.2 |  |  | 16
| 1 | 4 | 0.16 | 0.21 | 0.2 | 0.19 | < 1 | 0.26 | 0.22 | 0.18023 | 0.2 | 0.18 | 0.18 | 0.17 | 0.2 | 0.22611 | 0.2 | 0.19 | 0.19 | 0.2 | 0.2 | < 1 | 0.185 | 0.21 | 0.14 | 0.19 |  |  | 0.19621
| 1 | 5 | 0.16 | 0.2 | 0.27 | 0.19 | < 1 | 0.22 | 0.21 | 0.18687 | 0.2 | 0.19 | 0.2 | 0.17 | 0.19 | 0.22481 | 0.19 | 0.19 | 0.17 | 0.2 | 0.2 | < 1 | 0.204 | 0.2 | 0.15 | 0.19 |  |  | 157
| 1 | 6 | 0.21 | 0.18 | 0.17 | 0.2 | < 1 | 0.23 | 0.2 | 0.21525 | 0.19 | 0.19 | 0.19 | 0.17 | 0.2 | 0.21722 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | < 1 | 0.269 | 0.21 | 0.14 | 0.2
| Mean |  | 0.19167 | 0.21 | 0.21 | 0.19333 |  | 0.23667 | 0.21 | 0.19501 | 0.2 | 0.19167 | 0.18833 | 0.175 | 0.195 | 0.22603 | 0.20167 | 0.19167 | 0.198 | 0.2 | 0.2 |  | 0.21817 | 0.19833 | 0.135 | 0.19167
| Median |  | 0.2 | 0.205 | 0.2 | 0.19 |  | 0.235 | 0.21 | 0.18749 | 0.2 | 0.19 | 0.19 | 0.175 | 0.195 | 0.22568 | 0.2 | 0.19 | 0.2 | 0.2 | 0.2 |  | 0.2155 | 0.2 | 0.135 | 0.19
| Std Dev. |  | 0.025626 | 0.026833 | 0.041473 | 0.005164 |  | 0.01633 | 0.0089443 | 0.017863 | 0.0063246 | 0.0075277 | 0.0075277 | 0.0054772 | 0.010488 | 0.0058091 | 0.01169 | 0.0040825 | 0.021679 | 3.0405e-17 | 3.0405e-17 |  | 0.027953 | 0.01169 | 0.010488 | 0.0075277
| Rel.Std.Dev. |  | 13.37% | 12.78% | 19.75% | 2.671% |  | 6.9% | 4.259% | 9.16% | 3.162% | 3.928% | 3.997% | 3.13% | 5.379% | 2.57% | 5.797% | 2.13% | 10.95% | 1.52e-14% | 1.52e-14% |  | 12.81% | 5.894% | 7.769% | 3.928%
| PDM3 |  | -2.315% | 7.029% | 7.029% | -1.466% |  | 20.62% | 7.029% | -0.6088% | 1.932% | -2.315% | -4.014% | -10.81% | -0.6163% | 15.2% | 2.781% | -2.315% | 0.9127% | 1.932% | 1.932% |  | 11.19% | 1.083% | -31.2% | -2.315%
| Z-Score (Absolute) |  | 0.54 | 0.98 | 0.98 | 0.4 | 24.92 | 3.18 | 0.98 | 0.26 | 0.15 | 0.54 | 0.81 | 1.91 | 0.26 | 2.3 | 0.29 | 0.54 | 0.01 | NA | NA | 24.92 | 1.65 | 0.01 | 5.22 | 0.54
| Analytical results for Ce in OREAS 550b (Certified Value 43.3 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 1
| 1 | 1 | 44.366 | 47.8 | 55.5 | 91.32 | 40.56 | 34.67 | 43.3022 | 42.6 | 42.2 | 40.1 | 49.2 | 42.1 | 49.7729 | 36 | 41.64 | 67.35 | 42 | 38.9 | 48.342 | 69.6 | 80.8 | 36.74 |  |  |  |  | 1
| 1 | 2 | 44.116 | 48.6 | 55.2 | 91.896 | 41.41 | 36.36 | 41.4386 | 44.1 | 44 | 40.1 | 48.9 | 42.6 | 48.0374 | 36.8 | 42.64 | 64.03 | 44 | 39.8 | 48.222 | 79.9 | 75.9 | 37.29 |  |  |  |  | 9
| 1 | 3 | 44.472 | 47.6 | 56.8 | 91.232 | 41.33 | 34.25 | 43.1013 | 42.5 | 44.6 | 42 | 48.9 | 41.1 | 47.864 | 37.3 | 42.85 | 63.5 | 41 | 39.5 | 47.91 | 86.9 | 80.2 | 38.83 |  |  |  |  | 16
| 1 | 4 | 45.208 | 48.4 | 56.4 | 91.896 | 40.4 | 35.64 | 43.6995 | 43.8 | 43.5 | 40 | 47 | 43.8 | 47.9874 | 37 | 41.49 | 65.06 | 40 | 38.2 | 49.42 | 99.7 | 81.3 | 36.78 |  |  |  |  | 43.3083
| 1 | 5 | 43.924 | 46.2 | 56.1 | 91.128 | 42.76 | 36 | 43.2393 | 42.5 | 42.9 | 41 | 47.8 | 43.6 | 47.0644 | 36.8 | 43.11 | 62.67 | 41 | 38.4 | 49.416 | 52.4 | 78.6 | 36.09 |  |  |  |  | 158
| 1 | 6 | 44.234 | 46.9 | 57.5 | 91.112 | 41.31 | 34.25 | 41.3306 | 43.8 | 43.8 | 40.8 | 47.9 | 43.7 | 48.9555 | 36.7 | 42.76 | 62.59 | 41 | 38.9 | 48.276 | 100 | 84.2 | 37.24
| Mean |  | 44.3867 | 47.5833 | 56.25 | 91.4307 | 41.295 | 35.195 | 42.6853 | 43.2167 | 43.5 | 40.6667 | 48.2833 | 42.8167 | 48.2802 | 36.7667 | 42.415 | 64.2 | 41.5 | 38.95 | 48.5977 | 81.4167 | 80.1667 | 37.1617
| Median |  | 44.3 | 47.7 | 56.25 | 91.276 | 41.32 | 35.155 | 43.1703 | 43.2 | 43.65 | 40.45 | 48.4 | 43.1 | 48.0124 | 36.8 | 42.7 | 63.765 | 41 | 38.9 | 48.309 | 83.4 | 80.5 | 37.01
| Std Dev. |  | 0.44563 | 0.90866 | 0.84558 | 0.36822 | 0.83692 | 0.92357 | 1.0275 | 0.75741 | 0.84853 | 0.77374 | 0.85186 | 1.0834 | 0.94663 | 0.43205 | 0.67796 | 1.7952 | 1.3784 | 0.61563 | 0.65249 | 18.4043 | 2.779 | 0.92497
| Rel.Std.Dev. |  | 1.004% | 1.91% | 1.503% | 0.4027% | 2.027% | 2.624% | 2.407% | 1.753% | 1.951% | 1.903% | 1.764% | 2.53% | 1.961% | 1.175% | 1.598% | 2.796% | 3.321% | 1.581% | 1.343% | 22.61% | 3.466% | 2.489%
| PDM3 |  | 2.49% | 9.871% | 29.88% | 111.1% | -4.649% | -18.73% | -1.439% | -0.2116% | 0.4426% | -6.1% | 11.49% | -1.135% | 11.48% | -15.1% | -2.063% | 48.24% | -4.175% | -10.06% | 12.21% | 87.99% | 85.11% | -14.19%
| Z-Score (Absolute) |  | 0.15 | 0.61 | 1.86 | 6.95 | 0.3 | 1.18 | 0.1 | 0.02 | 0.02 | 0.39 | 0.71 | 0.08 | 0.71 | 0.95 | 0.14 | 3.01 | 0.27 | 0.64 | 0.76 | 5.5 | 5.32 | 0.9
| Analytical results for Co in OREAS 550b (Certified Value 0.131 wt.%) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 0.1335 | 0.11642 | 0.152 | 0.129 | 0.1214 | 0.1867 | 0.1325 | 0.14841 | 0.1447 | 0.1245 | 0.1155 | 0.1245 | 0.1345 | 0.121 | 0.13591 | 0.1435 | 0.1382 | 0.1472 | 0.1405 | 0.127 | 0.12782 | 0.1031 | 0.13876 | 0.119 | 0.139 | 0.1085 | 1
| 1 | 2 | 0.1323 | 0.12142 | 0.139 | 0.133 | 0.12074 | 0.1788 | 0.1335 | 0.15163 | 0.14296 | 0.125 | 0.12 | 0.1225 | 0.136 | 0.118 | 0.12929 | 0.1445 | 0.1404 | 0.1414 | 0.1425 | 0.132 | 0.12977 | 0.1023 | 0.13815 | 0.119 | 0.142 | 0.1096 | 29
| 1 | 3 | 0.13255 | 0.12589 | 0.139 | 0.136 | 0.12063 | 0.1676 | 0.1338 | 0.1511 | 0.14456 | 0.1235 | 0.1205 | 0.124 | 0.133 | 0.119 | 0.13289 | 0.1435 | 0.1364 | 0.143 | 0.1423 | 0.134 | 0.13095 | 0.1019 | 0.13856 | 0.12 | 0.134 | 0.1107 | 16
| 1 | 4 | 0.13499 | 0.12699 | 0.142 | 0.135 | 0.12081 | 0.1652 | 0.1332 | 0.14852 | 0.14391 | 0.1245 | 0.121 | 0.1195 | 0.1335 | 0.1185 | 0.1325 | 0.142 | 0.134 | 0.1474 | 0.1412 | 0.13 | 0.12718 | 0.1024 | 0.14359 | 0.126 | 0.135 | 0.1089 | 0.13079
| 1 | 5 | 0.13421 | 0.11381 | 0.138 | 0.134 | 0.12055 | 0.1785 | 0.1329 | 0.1484 | 0.13858 | 0.1235 | 0.1185 | 0.1205 | 0.133 | 0.119 | 0.13097 | 0.1425 | 0.1375 | 0.1459 | 0.1423 | 0.135 | 0.13104 | 0.1027 | 0.14155 | 0.128 | 0.14 | 0.1095 | 159
| 1 | 6 | 0.13024 | 0.12408 | 0.153 | 0.135 | 0.12186 | 0.1678 | 0.1354 | 0.15151 | 0.13938 | 0.1235 | 0.121 | 0.1205 | 0.1325 | 0.119 | 0.13005 | 0.1425 | 0.1363 | 0.1444 | 0.1404 | 0.131 | 0.12868 | 0.1023 | 0.14074 | 0.126 | 0.14 | 0.1094
| Mean |  | 0.13297 | 0.12144 | 0.14383 | 0.13367 | 0.121 | 0.1741 | 0.13355 | 0.14993 | 0.14235 | 0.12408 | 0.11942 | 0.12192 | 0.13375 | 0.11908 | 0.13194 | 0.14308 | 0.13713 | 0.14488 | 0.14153 | 0.1315 | 0.12924 | 0.10245 | 0.14022 | 0.123 | 0.13833 | 0.10943
| Median |  | 0.13303 | 0.12275 | 0.1405 | 0.1345 | 0.12077 | 0.17315 | 0.13335 | 0.14981 | 0.14344 | 0.124 | 0.12025 | 0.1215 | 0.13325 | 0.119 | 0.13173 | 0.143 | 0.13695 | 0.14515 | 0.14175 | 0.1315 | 0.12923 | 0.10235 | 0.13975 | 0.123 | 0.1395 | 0.10945
| Std Dev. |  | 0.0016719 | 0.0053102 | 0.0068532 | 0.0025033 | 0.00051846 | 0.0085013 | 0.0010134 | 0.0016367 | 0.0026933 | 0.00066458 | 0.0021311 | 0.0020595 | 0.0012942 | 0.0010206 | 0.0023889 | 0.00091742 | 0.0021445 | 0.0023937 | 0.00095638 | 0.002881 | 0.0016136 | 0.00040866 | 0.0021253 | 0.0040988 | 0.0031411 | 0.00074744
| Rel.Std.Dev. |  | 1.257% | 4.373% | 4.765% | 1.873% | 0.4285% | 4.883% | 0.7588% | 1.092% | 1.892% | 0.5356% | 1.785% | 1.689% | 0.9676% | 0.8571% | 1.811% | 0.6412% | 1.564% | 1.652% | 0.6757% | 2.191% | 1.249% | 0.3989% | 1.516% | 3.332% | 2.271% | 0.683%
| PDM3 |  | 1.664% | -7.152% | 9.974% | 2.2% | -7.486% | 33.12% | 2.111% | 14.63% | 8.837% | -5.127% | -8.695% | -6.784% | 2.264% | -8.95% | 0.8772% | 9.4% | 4.851% | 10.78% | 8.215% | 0.5437% | -1.184% | -21.67% | 7.214% | -5.955% | 5.768% | -16.33%
| Z-Score (Absolute) |  | 0.02 | 0.84 | 0.75 | 0.03 | 0.87 | 2.9 | 0.02 | 1.18 | 0.65 | 0.65 | 0.98 | 0.8 | 0.03 | 1.01 | 0.09 | 0.7 | 0.28 | 0.83 | 0.59 | 0.12 | 0.29 | 2.19 | 0.49 | 0.73 | 0.36 | 1.69
| Analytical results for Cr in OREAS 550b (Certified Value 28.5 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 20 | 21 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 2
| 1 | 1 | 30 | 27.3 | 31 | 32 | 42 | 30 | 24 | 28.9043 | 28 | 27 | 27 | 30 | 27 | 30.5543 | 25.1 | 29 | 29 | 29 | 29.044 | 32 | 41 | 28 |  |  |  |  | 1
| 1 | 2 | 30 | 27.8 | 31 | 32 | 43 | 30 | 26 | 27.3631 | 28 | 27 | 26 | 29 | 26 | 28.7979 | 25.6 | 30 | 29 | 29 | 28.773 | 7 | 40 | 27 |  |  |  |  | #N/A
| 1 | 3 | 29 | 28.8 | 31 | 33 | 37 | 30 | 25 | 27.9916 | 28 | 27 | 26 | 29 | 26 | 29.8402 | 26 | 29 | 27 | 30 | 28.794 | < 1 | 44 | 28 |  |  |  |  | 16
| 1 | 4 | 29 | 28.7 | 31 | 33 | 39 | 30 | 27 | 28.3326 | 28 | 27 | 25 | 29 | 26 | 29.6862 | 25.4 | 30 | 29 | 29 | 29.653 | 45 | 37 | 27 |  |  |  |  | 28.4997
| 1 | 5 | 30 | 26.3 | 31 | 33 | 41 | 30 | 26 | 29.2722 | 28 | 27 | 26 | 28 | 26 | 29.8215 | 27.1 | 29 | 30 | 30 | 29.615 | 48 | 37 | 28 |  |  |  |  | 160
| 1 | 6 | 30 | 28.1 | 33 | 34 | 34 | 29 | 25 | 26.9418 | 28 | 27 | 26 | 28 | 26 | 28.9515 | 25.3 | 30 | 31 | 30 | 29.129 | 56 | 44 | 28
| Mean |  | 29.6667 | 27.8333 | 31.3333 | 32.8333 | 39.3333 | 29.8333 | 25.5 | 28.1343 | 28 | 27 | 26 | 28.8333 | 26.1667 | 29.6086 | 25.75 | 29.5 | 29.1667 | 29.5 | 29.168 | 37.6 | 40.5 | 27.6667
| Median |  | 30 | 27.95 | 31 | 33 | 40 | 30 | 25.5 | 28.1621 | 28 | 27 | 26 | 29 | 26 | 29.7539 | 25.5 | 29.5 | 29 | 29.5 | 29.0865 | 45 | 40.5 | 28
| Std Dev. |  | 0.5164 | 0.93737 | 0.8165 | 0.75277 | 3.3862 | 0.40825 | 1.0488 | 0.8904 | 0 | 0 | 0.63246 | 0.75277 | 0.40825 | 0.64627 | 0.7287 | 0.54772 | 1.3292 | 0.54772 | 0.38674 | 19.1651 | 3.1464 | 0.5164
| Rel.Std.Dev. |  | 1.741% | 3.368% | 2.606% | 2.293% | 8.609% | 1.368% | 4.113% | 3.165% | 0% | 0% | 2.433% | 2.611% | 1.56% | 2.183% | 2.83% | 1.857% | 4.557% | 1.857% | 1.326% | 50.97% | 7.769% | 1.866%
| PDM3 |  | 4.095% | -2.338% | 9.943% | 15.21% | 38.01% | 4.679% | -10.53% | -1.282% | -1.753% | -5.262% | -8.771% | 1.171% | -8.186% | 3.891% | -9.648% | 3.51% | 2.34% | 3.51% | 2.345% | 31.93% | 42.11% | -2.923%
| Z-Score (Absolute) |  | 0.24 | 0.63 | 1.03 | 1.74 | 4.84 | 0.32 | 1.74 | 0.49 | 0.56 | 1.03 | 1.51 | 0.16 | 1.43 | 0.21 | 1.63 | 0.16 | 0 | 0.16 | 0 | 1.07 | 5.39 | 0.71
| Analytical results for Cs in OREAS 550b (Certified Value 0.45 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 2
| 1 | 1 | 0.58 | 0.46 | 0.72 | 1.3105 | 0.52 | 0.42 | 0.47936 | 0.4 | 0.42 | 0.39 | 0.44 | 0.44 | 0.54317 | 0.46 | 0.39 | 0.3 | 0.726 | 0.23 | 1 | 0.5 | [6 empty] | 1
| 1 | 2 | 0.59 | 0.53 | 0.74 | 1.3085 | 0.55 | 0.44 | 0.4255 | 0.42 | 0.44 | 0.4 | 0.44 | 0.44 | 0.57045 | 0.45 | 0.42 | 0.3 | 0.662 | 0.24 | 0.9 | 0.5 | [6 empty] | 12
| 1 | 3 | 0.58 | 0.49 | 0.71 | 1.2783 | 0.54 | 0.4 | 0.42838 | 0.41 | 0.44 | 0.4 | 0.44 | 0.45 | 0.56898 | 0.49 | 0.41 | 0.3 | 0.626 | 0.24 | 0.9 | 0.51 | [6 empty] | 16
| 1 | 4 | 0.59 | 0.56 | 0.7 | 1.3 | 0.55 | 0.45 | 0.44639 | 0.41 | 0.42 | 0.4 | 0.43 | 0.45 | 0.54575 | 0.45 | 0.39 | 0.3 | 0.639 | 0.25 | 0.7 | 0.5 | [6 empty] | 0.4544
| 1 | 5 | 0.59 | 0.46 | 0.73 | 1.267 | 0.54 | 0.41 | 0.46213 | 0.39 | 0.41 | 0.38 | 0.43 | 0.45 | 0.53098 | 0.47 | 0.4 | 0.3 | 0.633 | 0.25 | 0.8 | 0.49 | [6 empty] | 161
| 1 | 6 | 0.56 | 0.48 | 0.72 | 1.2849 | 0.53 | 0.43 | 0.46067 | 0.4 | 0.42 | 0.4 | 0.44 | 0.44 | 0.56424 | 0.46 | 0.39 | 0.3 | 0.614 | 0.24 | 1 | 0.5
| Mean |  | 0.58167 | 0.49667 | 0.72 | 1.2915 | 0.53833 | 0.425 | 0.4504 | 0.405 | 0.425 | 0.395 | 0.43667 | 0.445 | 0.55393 | 0.46333 | 0.4 | 0.3 | 0.65 | 0.24167 | 0.88333 | 0.5
| Median |  | 0.585 | 0.485 | 0.72 | 1.2925 | 0.54 | 0.425 | 0.45353 | 0.405 | 0.42 | 0.4 | 0.44 | 0.445 | 0.555 | 0.46 | 0.395 | 0.3 | 0.636 | 0.24 | 0.9 | 0.5
| Std Dev. |  | 0.01169 | 0.040332 | 0.014142 | 0.017548 | 0.01169 | 0.018708 | 0.020988 | 0.010488 | 0.012247 | 0.0083666 | 0.005164 | 0.0054772 | 0.016218 | 0.015055 | 0.012649 | 0 | 0.040502 | 0.0075277 | 0.1169 | 0.0063246
| Rel.Std.Dev. |  | 2.01% | 8.121% | 1.964% | 1.359% | 2.172% | 4.402% | 4.66% | 2.59% | 2.882% | 2.118% | 1.183% | 1.231% | 2.928% | 3.249% | 3.162% | 0% | 6.231% | 3.115% | 13.23% | 1.265%
| PDM3 |  | 28.01% | 9.302% | 58.45% | 184.2% | 18.47% | -6.47% | -0.8795% | -10.87% | -6.47% | -13.07% | -3.903% | -2.069% | 21.9% | 1.966% | -11.97% | -33.98% | 43.05% | -46.82% | 94.4% | 10.04%
| Z-Score (Absolute) |  | 1.42 | 0.45 | 2.99 | 9.48 | 0.93 | 0.36 | 0.07 | 0.59 | 0.36 | 0.7 | 0.23 | 0.13 | 1.1 | 0.07 | 0.65 | 1.78 | 2.19 | 2.44 | 4.84 | 0.49
| Analytical results for Cu in OREAS 550b (Certified Value 1.82 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*MS | AR*AAS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*AAS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*AAS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g |  | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 1.81 | 1.812 | 1.81 | 1.72 | 1.7378 | > 1 | 1.83 | 1.83 | 1.8331 | 1.84 | 1.865 | 1.82 | 1.87 | 1.785 | 1.8849 | 1.814 | 1.8 | 1.8406 | 1.8525 | 1.761 | > 1 | 1.82 | 1.762 | 1.8 | 1.8145 | 1.7 | 1
| 1 | 2 | 1.7843 | 1.814 | 1.82 | 1.73 | 1.7275 | > 1 | 1.83 | 1.83 | 1.7725 | 1.835 | 1.865 | 1.815 | 1.865 | 1.815 | 1.7769 | 1.822 | 1.84 | 1.7899 | 1.8738 | 1.794 | > 1 | 1.81 | 1.7511 | 1.78 | 1.8098 | 1.73 | 30
| 1 | 3 | 1.79 | 1.776 | 1.8 | 1.76 | 1.7256 | > 1 | 1.83 | 1.84 | 1.8068 | 1.85 | 1.875 | 1.84 | 1.855 | 1.78 | 1.8488 | 1.799 | 1.8 | 1.7985 | 1.858 | 1.79 | > 1 | 1.8 | 1.7639 | 1.82 | 1.8166 | 1.7 | 16
| 1 | 4 | 1.8242 | 1.825 | 1.81 | 1.77 | 1.7286 | > 1 | 1.84 | 1.84 | 1.8188 | 1.825 | 1.85 | 1.84 | 1.865 | 1.82 | 1.8378 | 1.802 | 1.83 | 1.8609 | 1.8636 | 1.778 | > 1 | 1.79 | 1.8399 | 1.86 | 1.8099 | 1.72 | 1.8171
| 1 | 5 | 1.816 | 1.805 | 1.79 | 1.75 | 1.7256 | > 1 | 1.85 | 1.83 | 1.7963 | 1.855 | 1.84 | 1.835 | 1.88 | 1.78 | 1.8052 | 1.817 | 1.83 | 1.8401 | 1.8761 | 1.794 | > 1 | 1.8 | 1.82 | 1.85 | 1.812 | 1.7 | 162
| 1 | 6 | 1.7591 | 1.806 | 1.8 | 1.78 | 1.7302 | > 1 | 1.85 | 1.84 | 1.7585 | 1.845 | 1.88 | 1.83 | 1.855 | 1.755 | 1.8135 | 1.793 | 1.83 | 1.8028 | 1.8658 | 1.79 | > 1 | 1.79 | 1.7998 | 1.83 | 1.8172 | 1.71
| Mean |  | 1.7973 | 1.8063 | 1.805 | 1.7517 | 1.7292 |  | 1.8383 | 1.835 | 1.7977 | 1.8417 | 1.8625 | 1.83 | 1.865 | 1.7892 | 1.8279 | 1.8078 | 1.8217 | 1.8221 | 1.865 | 1.7845 |  | 1.8017 | 1.7895 | 1.8233 | 1.8133 | 1.71
| Median |  | 1.8 | 1.809 | 1.805 | 1.755 | 1.728 |  | 1.835 | 1.835 | 1.8016 | 1.8425 | 1.865 | 1.8325 | 1.865 | 1.7825 | 1.8256 | 1.808 | 1.83 | 1.8215 | 1.8647 | 1.79 |  | 1.8 | 1.7818 | 1.825 | 1.8133 | 1.705
| Std Dev. |  | 0.024136 | 0.016501 | 0.010488 | 0.023166 | 0.0045635 |  | 0.0098319 | 0.0054772 | 0.02815 | 0.010801 | 0.015083 | 0.010488 | 0.0094868 | 0.024376 | 0.037708 | 0.011444 | 0.017224 | 0.028768 | 0.0090387 | 0.012927 |  | 0.01169 | 0.035941 | 0.030111 | 0.0032504 | 0.012649
| Rel.Std.Dev. |  | 1.343% | 0.9135% | 0.5811% | 1.323% | 0.2639% |  | 0.5348% | 0.2985% | 1.566% | 0.5865% | 0.8098% | 0.5731% | 0.5087% | 1.362% | 2.063% | 0.633% | 0.9455% | 1.579% | 0.4847% | 0.7244% |  | 0.6489% | 2.008% | 1.651% | 0.1793% | 0.7397%
| PDM3 |  | -1.092% | -0.5929% | -0.6663% | -3.601% | -4.837% |  | 1.168% | 0.9847% | -1.069% | 1.352% | 2.498% | 0.7095% | 2.636% | -1.538% | 0.5916% | -0.5104% | 0.2509% | 0.2766% | 2.634% | -1.794% |  | -0.8497% | -1.522% | 0.3427% | -0.2076% | -5.894%
| Z-Score (Absolute) |  | 0.44 | 0.14 | 0.19 | 1.96 | 2.71 | NA | 0.92 | 0.81 | 0.43 | 1.03 | 1.73 | 0.65 | 1.81 | 0.71 | 0.57 | 0.09 | 0.37 | 0.38 | 1.81 | 0.87 | NA | 0.3 | 0.7 | 0.42 | 0.09 | 3.35
| Analytical results for Dy in OREAS 550b (Certified Value 1.17 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 3
| 1 | 1 | 1.204 | 1.2 | 2.2766 | 1.3925 | 1.148 | 0.95 | 1.13 | [19 empty] | 1
| 1 | 2 | 1.202 | 1.2 | 2.305 | 1.3233 | 1.135 | 0.94 | 1.1 | [19 empty] | 13
| 1 | 3 | 1.224 | 1.2 | 2.2918 | 1.3579 | 1.14 | 0.98 | 1.23 | [19 empty] | 16
| 1 | 4 | 1.24 | 1.2 | 2.315 | 1.3223 | 1.187 | 0.98 | 1.18 | [19 empty] | 1.1679
| 1 | 5 | 1.216 | 1.1 | 2.3307 | 1.3159 | 1.187 | 1 | 1.1 | [19 empty] | 163
| 1 | 6 | 1.274 | 1.1 | 2.2881 | 1.3156 | 1.149 | 0.97 | 1.15
| Mean |  | 1.2267 | 1.1667 | 2.3012 | 1.3379 | 1.1577 | 0.97 | 1.1483
| Median |  | 1.22 | 1.2 | 2.2984 | 1.3228 | 1.1485 | 0.975 | 1.14
| Std Dev. |  | 0.02706 | 0.05164 | 0.019707 | 0.031048 | 0.023304 | 0.021909 | 0.050365
| Rel.Std.Dev. |  | 2.206% | 4.426% | 0.8564% | 2.321% | 2.013% | 2.259% | 4.386%
| PDM3 |  | 5.034% | -0.1032% | 97.04% | 14.56% | -0.8739% | -16.94% | -1.673%
| Z-Score (Absolute) |  | 0.67 | 0 | 12.75 | 1.92 | 0.1 | 2.21 | 0.21
| Analytical results for Er in OREAS 550b (Certified Value 0.52 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 3
| 1 | 1 | 0.527 | 0.6 | 1.0456 | 0.57455 | 0.5 | 0.45 | 0.55 | [19 empty] | 1
| 1 | 2 | 0.51 | 0.5 | 1.0302 | 0.57095 | 0.489 | 0.47 | 0.53 | [19 empty] | 14
| 1 | 3 | 0.513 | 0.6 | 1.0073 | 0.56336 | 0.476 | 0.46 | 0.55 | [19 empty] | 16
| 1 | 4 | 0.505 | 0.5 | 1.0348 | 0.55961 | 0.506 | 0.48 | 0.54 | [19 empty] | 0.51885
| 1 | 5 | 0.509 | 0.5 | 1.0604 | 0.56275 | 0.499 | 0.49 | 0.53 | [19 empty] | 164
| 1 | 6 | 0.504 | 0.5 | 1.0914 | 0.55022 | 0.499 | 0.46 | 0.55
| Mean |  | 0.51133 | 0.53333 | 1.0449 | 0.56357 | 0.49483 | 0.46833 | 0.54167
| Median |  | 0.5095 | 0.5 | 1.0402 | 0.56305 | 0.499 | 0.465 | 0.545
| Std Dev. |  | 0.0083586 | 0.05164 | 0.028745 | 0.0085973 | 0.010722 | 0.01472 | 0.0098319
| Rel.Std.Dev. |  | 1.635% | 9.682% | 2.751% | 1.526% | 2.167% | 3.143% | 1.815%
| PDM3 |  | -1.448% | 2.792% | 101.4% | 8.621% | -4.628% | -9.735% | 4.398%
| Z-Score (Absolute) |  | 0.49 | 0 | 11.41 | 0.67 | 0.86 | 1.45 | 0.19
| Analytical results for Eu in OREAS 550b (Certified Value 0.39 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 0.4 | 0.4 | 0.92328 | 0.46066 | 0.41 | 0.29 | 0.37 | [19 empty] | 1
| 1 | 2 | 0.409 | 0.4 | 0.90576 | 0.46304 | 0.401 | 0.29 | 0.36 | [19 empty] | 15
| 1 | 3 | 0.403 | 0.4 | 0.89464 | 0.4566 | 0.404 | 0.3 | 0.41 | [19 empty] | 16
| 1 | 4 | 0.418 | 0.4 | 0.88504 | 0.45846 | 0.415 | 0.3 | 0.39 | [19 empty] | 0.39055
| 1 | 5 | 0.408 | 0.4 | 0.8952 | 0.45434 | 0.404 | 0.31 | 0.36 | [19 empty] | 165
| 1 | 6 | 0.402 | 0.4 | 0.9204 | 0.43377 | 0.409 | 0.3 | 0.37
| Mean |  | 0.40667 | 0.4 | 0.90405 | 0.45448 | 0.40717 | 0.29833 | 0.37667
| Median |  | 0.4055 | 0.4 | 0.90048 | 0.45753 | 0.4065 | 0.3 | 0.37
| Std Dev. |  | 0.0065625 | 6.0809e-17 | 0.015286 | 0.010592 | 0.0051153 | 0.0075277 | 0.019664
| Rel.Std.Dev. |  | 1.614% | 1.52e-14% | 1.691% | 2.33% | 1.256% | 2.523% | 5.22%
| PDM3 |  | 4.126% | 2.419% | 131.5% | 16.37% | 4.254% | -23.61% | -3.555%
| Z-Score (Absolute) |  | 0 | 0.15 | 11.18 | 1.07 | 0.01 | 2.44 | 0.67
| Analytical results for Fe in OREAS 550b (Certified Value 2.46 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 2.54 | 2.4 | 2.44 | 2.44 | 2.4347 | 3.03 | 2.45 | 2.48 | 2.4253 | 2.52 | 2.45 | 2.5 | 2.53 | 2.45 | 2.5057 | 2.44 | 2.44 | 2.59 | 2.3494 | 2.3 | 2.45 | 2.34 | 2.3634 | 2.868 | 2.48 | 2.45 | 1
| 1 | 2 | 2.57 | 2.37 | 2.47 | 2.51 | 2.4335 | 3.2 | 2.45 | 2.59 | 2.3905 | 2.54 | 2.49 | 2.47 | 2.57 | 2.39 | 2.4943 | 2.49 | 2.45 | 2.5 | 2.4233 | 2.3 | 2.48 | 2.33 | 2.3389 | 2.819 | 2.58 | 2.46 | #N/A
| 1 | 3 | 2.45 | 2.42 | 2.49 | 2.54 | 2.4245 | 2.73 | 2.45 | 2.47 | 2.4818 | 2.53 | 2.49 | 2.47 | 2.53 | 2.4 | 2.5001 | 2.46 | 2.46 | 2.5 | 2.4082 | 2.3 | 2.47 | 2.35 | 2.3409 | 2.868 | 2.65 | 2.45 | 16
| 1 | 4 | 2.42 | 2.43 | 2.44 | 2.46 | 2.4321 | 2.66 | 2.46 | 2.63 | 2.5022 | 2.53 | 2.51 | 2.4 | 2.53 | 2.39 | 2.5462 | 2.51 | 2.45 | 2.6 | 2.4166 | 2.35 | 2.46 | 2.33 | 2.4518 | 2.882 | 2.55 | 2.44 | 2.4576
| 1 | 5 | 2.57 | 2.33 | 2.48 | 2.51 | 2.4361 | 3.04 | 2.46 | 2.46 | 2.5424 | 2.52 | 2.44 | 2.42 | 2.53 | 2.4 | 2.4866 | 2.51 | 2.42 | 2.55 | 2.3932 | 2.31 | 2.47 | 2.34 | 2.4057 | 2.868 | 2.51 | 2.43 | 166
| 1 | 6 | 2.52 | 2.24 | 2.4 | 2.49 | 2.4273 | 2.65 | 2.45 | 2.54 | 2.3565 | 2.5 | 2.51 | 2.41 | 2.51 | 2.4 | 2.467 | 2.49 | 2.44 | 2.55 | 2.4015 | 2.38 | 2.45 | 2.35 | 2.379 | 2.889 | 2.56 | 2.47
| Mean |  | 2.5117 | 2.365 | 2.4533 | 2.4917 | 2.4314 | 2.885 | 2.4533 | 2.5283 | 2.4498 | 2.5233 | 2.4817 | 2.445 | 2.5333 | 2.405 | 2.5 | 2.4833 | 2.4433 | 2.5483 | 2.3987 | 2.3233 | 2.4633 | 2.34 | 2.38 | 2.8657 | 2.555 | 2.45
| Median |  | 2.53 | 2.385 | 2.455 | 2.5 | 2.4328 | 2.88 | 2.45 | 2.51 | 2.4535 | 2.525 | 2.49 | 2.445 | 2.53 | 2.4 | 2.4972 | 2.49 | 2.445 | 2.55 | 2.4048 | 2.305 | 2.465 | 2.34 | 2.3712 | 2.868 | 2.555 | 2.45
| Std Dev. |  | 0.063061 | 0.071204 | 0.033267 | 0.03656 | 0.0045259 | 0.23416 | 0.005164 | 0.070261 | 0.070923 | 0.013663 | 0.029944 | 0.040373 | 0.019664 | 0.022583 | 0.026344 | 0.028048 | 0.013663 | 0.042622 | 0.026398 | 0.033862 | 0.012111 | 0.0089443 | 0.043155 | 0.024517 | 0.058907 | 0.014142
| Rel.Std.Dev. |  | 2.511% | 3.011% | 1.356% | 1.467% | 0.1861% | 8.116% | 0.2105% | 2.779% | 2.895% | 0.5415% | 1.207% | 1.651% | 0.7762% | 0.939% | 1.054% | 1.129% | 0.5592% | 1.673% | 1.101% | 1.457% | 0.4916% | 0.3822% | 1.813% | 0.8555% | 2.306% | 0.5772%
| PDM3 |  | 2.199% | -3.769% | -0.1749% | 1.385% | -1.069% | 17.39% | -0.1749% | 2.877% | -0.3193% | 2.673% | 0.978% | -0.5139% | 3.08% | -2.142% | 1.724% | 1.046% | -0.5817% | 3.691% | -2.398% | -5.464% | 0.232% | -4.786% | -3.16% | 16.6% | 3.962% | -0.3105%
| Z-Score (Absolute) |  | 0.67 | 1.18 | 0.06 | 0.42 | 0.34 | 5.39 | 0.06 | 0.89 | 0.11 | 0.82 | 0.3 | 0.17 | 0.95 | 0.67 | 0.53 | 0.32 | 0.19 | 1.14 | 0.75 | 1.71 | 0.06 | 1.5 | 0.99 | 5.15 | 1.22 | 0.11
| Analytical results for Ga in OREAS 550b (Certified Value 7.31 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 2
| 1 | 1 | 7.18 | 6.6 | 7.86 | 9.8 | 7 | 6.2 | 7.7672 | 7.09 | 6.41 | 6.74 | 7.42 | 6.51 | 7.6691 | 6.73 | 7.4 | 6 | 8 | 8.2 | 7.667 | 4.4 | 9.99 | 6.51 |  |  |  |  | 1
| 1 | 2 | 7.19 | 6.8 | 8.14 | 10 | 8 | 6.4 | 7.1217 | 7.13 | 6.71 | 6.93 | 7.95 | 6.55 | 7.7417 | 6.76 | 7.6 | 6 | 8 | 8.2 | 7.712 | 4.29 | 9.41 | 6.32 |  |  |  |  | 17
| 1 | 3 | 7.24 | 7.3 | 8.3 | 10 | 8 | 6.7 | 7.4522 | 6.8 | 6.73 | 7.09 | 8.01 | 6.59 | 7.8578 | 6.87 | 7.6 | 6 | 8 | 8.2 | 7.729 | 4.44 | 10.1 | 6.55 |  |  |  |  | 16
| 1 | 4 | 7.2 | 7.6 | 8.35 | 9.9 | 8 | 6.1 | 7.5205 | 7.1 | 6.59 | 6.77 | 7.26 | 6.7 | 7.8844 | 6.72 | 7.6 | 6 | 8 | 8.8 | 7.904 | 4.56 | 9.26 | 6.29 |  |  |  |  | 7.3108
| 1 | 5 | 7.12 | 6.5 | 8.17 | 10.2 | 7 | 6.5 | 7.9234 | 6.84 | 6.37 | 7.09 | 7.83 | 6.72 | 7.825 | 6.84 | 7.4 | 6 | 8 | 8.6 | 7.822 | 4.62 | 8.36 | 6.29 |  |  |  |  | 167
| 1 | 6 | 7.08 | 6.5 | 8.47 | 10 | 7 | 6.7 | 7.1577 | 6.97 | 6.59 | 6.93 | 7.54 | 6.57 | 7.5954 | 6.64 | 7.8 | 7 | 8 | 8.5 | 7.728 | 4.51 | 9.96 | 6.01
| Mean |  | 7.1683 | 6.8833 | 8.215 | 9.9833 | 7.5 | 6.4333 | 7.4905 | 6.9883 | 6.5667 | 6.925 | 7.6683 | 6.6067 | 7.7622 | 6.76 | 7.5667 | 6.1667 | 8 | 8.4167 | 7.7603 | 4.47 | 9.5133 | 6.3283
| Median |  | 7.185 | 6.7 | 8.235 | 10 | 7.5 | 6.45 | 7.4864 | 7.03 | 6.59 | 6.93 | 7.685 | 6.58 | 7.7833 | 6.745 | 7.6 | 6 | 8 | 8.35 | 7.7285 | 4.475 | 9.685 | 6.305
| Std Dev. |  | 0.058109 | 0.46224 | 0.21173 | 0.13292 | 0.54772 | 0.25033 | 0.32034 | 0.1419 | 0.14935 | 0.15017 | 0.30564 | 0.084538 | 0.11393 | 0.084143 | 0.15055 | 0.40825 | 0 | 0.25626 | 0.086623 | 0.11866 | 0.65969 | 0.19312
| Rel.Std.Dev. |  | 0.8106% | 6.715% | 2.577% | 1.331% | 7.303% | 3.891% | 4.277% | 2.031% | 2.274% | 2.168% | 3.986% | 1.28% | 1.468% | 1.245% | 1.99% | 6.62% | 0% | 3.045% | 1.116% | 2.655% | 6.934% | 3.052%
| PDM3 |  | -1.949% | -5.847% | 12.37% | 36.56% | 2.588% | -12% | 2.457% | -4.411% | -10.18% | -5.277% | 4.89% | -9.632% | 6.175% | -7.534% | 3.5% | -15.65% | 9.427% | 15.13% | 6.149% | -38.86% | 30.13% | -13.44%
| Z-Score (Absolute) |  | 0.18 | 0.49 | 0.96 | 2.89 | 0.19 | 0.97 | 0.18 | 0.37 | 0.83 | 0.44 | 0.37 | 0.79 | 0.47 | 0.62 | 0.26 | 1.26 | 0.73 | 1.18 | 0.47 | 3.11 | 2.38 | 1.09
| Analytical results for Gd in OREAS 550b (Certified Value 2.19 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 3
| 1 | 1 | 1.964 | 2.2 | 5.2238 | 2.5307 | 2.209 | 1.45 | 2.03 | [19 empty] | 1
| 1 | 2 | 1.999 | 2.2 | 5.2159 | 2.4914 | 2.123 | 1.44 | 2.09 | [19 empty] | 18
| 1 | 3 | 2.038 | 2.3 | 5.3143 | 2.5299 | 2.138 | 1.5 | 2.12 | [19 empty] | 16
| 1 | 4 | 2.016 | 2.4 | 5.2056 | 2.4447 | 2.262 | 1.51 | 2.04 | [19 empty] | 2.1919
| 1 | 5 | 1.977 | 2.2 | 5.1941 | 2.3973 | 2.213 | 1.55 | 1.99 | [19 empty] | 168
| 1 | 6 | 1.994 | 2.2 | 5.1938 | 2.3969 | 2.153 | 1.51 | 2.11
| Mean |  | 1.998 | 2.25 | 5.2246 | 2.4651 | 2.183 | 1.4933 | 2.0633
| Median |  | 1.9965 | 2.2 | 5.2108 | 2.468 | 2.181 | 1.505 | 2.065
| Std Dev. |  | 0.026601 | 0.083666 | 0.04554 | 0.061404 | 0.053557 | 0.041312 | 0.051251
| Rel.Std.Dev. |  | 1.331% | 3.718% | 0.8716% | 2.491% | 2.453% | 2.766% | 2.484%
| PDM3 |  | -8.846% | 2.651% | 138.4% | 12.47% | -0.4057% | -31.87% | -5.865%
| Z-Score (Absolute) |  | 0.67 | 0.24 | 11.09 | 1.03 | 0 | 2.51 | 0.44
| Analytical results for Ge in OREAS 550b (Indicative Value 0.077 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [10 empty] | 3
| 1 | 1 | < 0.1 | 0.21 | 0.1 | 0.05 | 0.06 | 0.09 | 0.06 | 0.07 | < 0.2 | < 0.1 | 1 | < 5 | 0.217 | < 0.01 | 0.1 | 0.12 | [10 empty] | 1
| 1 | 2 | < 0.1 | 0.21 | 0.1 | < 0.05 | 0.05 | 0.09 | 0.07 | 0.06 | < 0.2 | < 0.1 | 1.1 | < 5 | 0.196 | < 0.01 | 0.09 | 0.11 | [10 empty] | 2
| 1 | 3 | < 0.1 | 0.2 | 0.1 | 0.05 | 0.05 | 0.07 | 0.07 | 0.07 | < 0.2 | < 0.1 | 1.1 | < 5 | 0.203 | < 0.01 | 0.09 | 0.11 | [10 empty] | 16
| 1 | 4 | < 0.1 | 0.21 | 0.1 | < 0.05 | 0.05 | 0.08 | 0.06 | 0.07 | < 0.2 | < 0.1 | 1 | < 5 | 0.21 | < 0.01 | 0.1 | 0.1 | [10 empty] | 0.077083
| 1 | 5 | < 0.1 | 0.19 | 0.1 | 0.05 | 0.06 | 0.09 | 0.07 | 0.06 | < 0.2 | < 0.1 | 1.1 | < 5 | 0.196 | < 0.01 | 0.07 | 0.09 | [10 empty] | 30
| 1 | 6 | < 0.1 | 0.21 | 0.1 | 0.05 | 0.06 | 0.09 | 0.07 | 0.06 | < 0.2 | < 0.1 | 1.1 | < 5 | 0.198 | < 0.01 | 0.1 | 0.09
| Mean |  |  | 0.205 | 0.1 | 0.05 | 0.055 | 0.085 | 0.066667 | 0.065 |  |  | 1.0667 |  | 0.20333 |  | 0.091667 | 0.10333
| Median |  |  | 0.21 | 0.1 | 0.05 | 0.055 | 0.09 | 0.07 | 0.065 |  |  | 1.1 |  | 0.2005 |  | 0.095 | 0.105
| Std Dev. |  |  | 0.0083666 | 1.5202e-17 | 0 | 0.0054772 | 0.0083666 | 0.005164 | 0.0054772 |  |  | 0.05164 |  | 0.0085713 |  | 0.01169 | 0.012111
| Rel.Std.Dev. |  |  | 4.081% | 1.52e-14% | 0% | 9.959% | 9.843% | 7.746% | 8.427% |  |  | 4.841% |  | 4.215% |  | 12.75% | 11.72%
| PDM3 |  |  | 165.9% | 29.73% | -35.14% | -28.65% | 10.27% | -13.51% | -15.68% |  |  | 1284% |  | 163.8% |  | 18.92% | 34.05%
| Z-Score (Absolute) |  | 0.72 | 2.2 | 0.22 | 0.88 | 0.63 | 0.06 | 0.41 | 0.44 | 0.22 | 0.72 | 18.41 | 45.38 | 2.16 | 1.57 | 0.06 | 0.28
| Analytical results for Hf in OREAS 550b (Certified Value 0.32 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 23 | 24 | 25 | 26 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [7 empty] | 2
| 1 | 1 | 0.31 | 0.1 | 0.3 | 0.25 | 0.39 | 0.31727 | 0.28 | 0.32 | 0.32 | 0.27 | 0.31 | 0.36119 | 0.3 | 0.33 | 0.3 | 0.413 | 0.24 | 0.43 | 0.34 | [7 empty] | 1
| 1 | 2 | 0.31 | 0.2 | 0.31 | 0.25 | 0.35 | 0.33551 | 0.28 | 0.32 | 0.31 | 0.28 | 0.31 | 0.35979 | 0.32 | 0.33 | 0.3 | 0.387 | 0.25 | 0.4 | 0.32 | [7 empty] | 22
| 1 | 3 | 0.29 | 0.2 | 0.33 | 0.28 | 0.37 | 0.32237 | 0.29 | 0.32 | 0.31 | 0.27 | 0.3 | 0.36524 | 0.31 | 0.34 | 0.3 | 0.365 | 0.24 | 0.44 | 0.32 | [7 empty] | 16
| 1 | 4 | 0.31 | 0.2 | 0.33 | 0.26 | 0.4 | 0.34663 | 0.29 | 0.32 | 0.3 | 0.26 | 0.32 | 0.36534 | 0.34 | 0.32 | 0.3 | 0.388 | 0.24 | 0.44 | 0.35 | [7 empty] | 0.32104
| 1 | 5 | 0.29 | 0.2 | 0.32 | 0.28 | 0.35 | 0.33832 | 0.3 | 0.32 | 0.32 | 0.28 | 0.32 | 0.36568 | 0.36 | 0.31 | 0.3 | 0.379 | 0.25 | 0.45 | 0.34 | [7 empty] | 169
| 1 | 6 | 0.3 | < 0.1 | 0.32 | 0.27 | 0.4 | 0.30913 | 0.3 | 0.32 | 0.33 | 0.29 | 0.3 | 0.34928 | 0.34 | 0.32 | 0.3 | 0.382 | 0.24 | 0.49 | 0.35
| Mean |  | 0.30167 | 0.18 | 0.31833 | 0.265 | 0.37667 | 0.3282 | 0.29 | 0.32 | 0.315 | 0.275 | 0.31 | 0.36109 | 0.32833 | 0.325 | 0.3 | 0.38567 | 0.24333 | 0.44167 | 0.33667
| Median |  | 0.305 | 0.2 | 0.32 | 0.265 | 0.38 | 0.32894 | 0.29 | 0.32 | 0.315 | 0.275 | 0.31 | 0.36321 | 0.33 | 0.325 | 0.3 | 0.3845 | 0.24 | 0.44 | 0.34
| Std Dev. |  | 0.0098319 | 0.044721 | 0.01169 | 0.013784 | 0.023381 | 0.014232 | 0.0089443 | 0 | 0.010488 | 0.010488 | 0.0089443 | 0.0062863 | 0.022286 | 0.010488 | 0 | 0.015744 | 0.005164 | 0.029269 | 0.013663
| Rel.Std.Dev. |  | 3.259% | 24.85% | 3.672% | 5.202% | 6.207% | 4.336% | 3.084% | 0% | 3.33% | 3.814% | 2.885% | 1.741% | 6.788% | 3.227% | 0% | 4.082% | 2.122% | 6.627% | 4.058%
| PDM3 |  | -6.034% | -43.93% | -0.8427% | -17.46% | 17.33% | 2.232% | -9.668% | -0.3236% | -1.881% | -14.34% | -3.438% | 12.47% | 2.272% | 1.234% | -6.553% | 20.13% | -24.2% | 37.57% | 4.868%
| Z-Score (Absolute) |  | 0.61 | 5.88 | 0 | 1.96 | 2.15 | 0.36 | 1.04 | 0.06 | 0.12 | 1.59 | 0.31 | 1.57 | 0.37 | 0.25 | 0.67 | 2.48 | 2.76 | 4.54 | 0.67
| Analytical results for Hg in OREAS 550b (Certified Value 0.047 ppm) | [27 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 3
| 1 | 1 | 0.022 | 0.03 | 0.05 | 0.06 | < 0.05 | 0.04 | 0.06 | 0.03 | 0.04 | 0.05 | < 1 | < 1 | 0.05 | 0.06 | 0.05 | 0.054 | < 5 | 0.05 | 0.04 | 0.07 | [6 empty] | 1
| 1 | 2 | 0.023 | 0.04 | 0.05 | 0.04 | < 0.05 | 0.04 | 0.05 | 0.03 | 0.04 | 0.04 | < 1 | < 1 | 0.05 | 0.05 | 0.04 | 0.063 | < 5 | 0.05 | 0.04 | 0.11 | [6 empty] | 1
| 1 | 3 | 0.027 | 0.04 | 0.05 | 0.06 | < 0.05 | 0.03 | 0.05 | 0.03 | 0.04 | 0.04 | < 1 | < 1 | 0.04 | 0.06 | 0.05 | 0.061 | < 5 | 0.05 | 0.04 | 0.07 | [6 empty] | 16
| 1 | 4 | 0.024 | 0.05 | 0.05 | 0.05 | < 0.05 | 0.04 | 0.06 | 0.05 | 0.04 | 0.05 | < 1 | < 1 | 0.04 | 0.06 | 0.05 | 0.064 | < 5 | 0.06 | 0.04 | 0.06 | [6 empty] | 0.047214
| 1 | 5 | 0.016 | 0.04 | 0.05 | 0.04 | < 0.05 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | < 1 | < 1 | 0.05 | 0.05 | 0.04 | 0.06 | < 5 | 0.06 | 0.04 | 0.07 | [6 empty] | 170
| 1 | 6 | 0.045 | 0.02 | 0.05 | 0.05 | < 0.05 | 0.04 | 0.05 | 0.06 | 0.05 | 0.05 | < 1 | < 1 | 0.04 | 0.07 | 0.04 | 0.064 | < 5 | 0.06 | 0.05 | 0.09
| Mean |  | 0.026167 | 0.036667 | 0.05 | 0.05 |  | 0.038333 | 0.053333 | 0.04 | 0.041667 | 0.045 |  |  | 0.045 | 0.058333 | 0.045 | 0.061 |  | 0.055 | 0.041667 | 0.078333
| Median |  | 0.0235 | 0.04 | 0.05 | 0.05 |  | 0.04 | 0.05 | 0.035 | 0.04 | 0.045 |  |  | 0.045 | 0.06 | 0.045 | 0.062 |  | 0.055 | 0.04 | 0.07
| Std Dev. |  | 0.0099079 | 0.010328 | 7.6012e-18 | 0.0089443 |  | 0.0040825 | 0.005164 | 0.012649 | 0.0040825 | 0.0054772 |  |  | 0.0054772 | 0.0075277 | 0.0054772 | 0.0037947 |  | 0.0054772 | 0.0040825 | 0.018348
| Rel.Std.Dev. |  | 37.86% | 28.17% | 1.52e-14% | 17.89% |  | 10.65% | 9.682% | 31.62% | 9.798% | 12.17% |  |  | 12.17% | 12.9% | 12.17% | 6.221% |  | 9.959% | 9.798% | 23.42%
| PDM3 |  | -44.58% | -22.34% | 5.9% | 5.9% |  | -18.81% | 12.96% | -15.28% | -11.75% | -4.69% |  |  | -4.69% | 23.55% | -4.69% | 29.2% |  | 16.49% | -11.75% | 65.91%
| Z-Score (Absolute) |  | 1.73 | 0.88 | 0.2 | 0.2 | 1.82 | 0.74 | 0.47 | 0.61 | 0.47 | 0.2 | 36.61 | 36.61 | 0.2 | 0.88 | 0.2 | 1.09 | 198.45 | 0.61 | 0.47 | 2.49
| Analytical results for Ho in OREAS 550b (Certified Value 0.2 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 3
| 1 | 1 | 0.193 | 0.2 | 0.38992 | 0.22412 | 1.148 | 0.17 | 0.19 | [19 empty] | 1
| 1 | 2 | 0.199 | 0.2 | 0.3712 | 0.22092 | 1.135 | 0.16 | 0.19 | [19 empty] | 5
| 1 | 3 | 0.194 | 0.2 | 0.37152 | 0.22275 | 1.14 | 0.17 | 0.19 | [19 empty] | 16
| 1 | 4 | 0.202 | 0.2 | 0.40904 | 0.21913 | 1.187 | 0.17 | 0.2 | [19 empty] | 0.19551
| 1 | 5 | 0.196 | 0.2 | 0.37752 | 0.21301 | 1.187 | 0.18 | 0.19 | [19 empty] | 171
| 1 | 6 | 0.189 | 0.2 | 0.3896 | 0.21232 | 1.149 | 0.17 | 0.2
| Mean |  | 0.1955 | 0.2 | 0.3848 | 0.21871 | 1.1577 | 0.17 | 0.19333
| Median |  | 0.195 | 0.2 | 0.38356 | 0.22003 | 1.1485 | 0.17 | 0.19
| Std Dev. |  | 0.0045935 | 3.0405e-17 | 0.014499 | 0.00498 | 0.023304 | 0.0063246 | 0.005164
| Rel.Std.Dev. |  | 2.35% | 1.52e-14% | 3.768% | 2.277% | 2.013% | 3.72% | 2.671%
| PDM3 |  | -0.004495% | 2.297% | 96.82% | 11.87% | 492.1% | -13.05% | -1.113%
| Z-Score (Absolute) |  | 0.16 | 0 | 6.66 | 0.67 | 34.51 | 1.08 | 0.24
| Analytical results for In in OREAS 550b (Certified Value 0.63 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 25 | 26 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.1g | 0.4g | [7 empty] | 3
| 1 | 1 | 0.63 | 0.63 | 0.74 | 0.626 | 0.62 | 0.68655 | 0.624 | 0.662 | 0.588 | 0.614 | 0.611 | 0.67459 | 0.57 | 0.6 | 0.62 | < 5 | 0.56 | 0.66 | 0.626 | [7 empty] | 1
| 1 | 2 | 0.6 | 0.6 | 0.74 | 0.639 | 0.58 | 0.66209 | 0.636 | 0.679 | 0.586 | 0.626 | 0.638 | 0.67031 | 0.59 | 0.63 | 0.63 | < 5 | 0.555 | 0.66 | 0.637 | [7 empty] | 6
| 1 | 3 | 0.61 | 0.6 | 0.73 | 0.674 | 0.63 | 0.66448 | 0.635 | 0.687 | 0.617 | 0.623 | 0.641 | 0.66096 | 0.6 | 0.61 | 0.62 | < 5 | 0.55 | 0.67 | 0.619 | [7 empty] | 16
| 1 | 4 | 0.61 | 0.61 | 0.74 | 0.655 | 0.59 | 0.70033 | 0.64 | 0.697 | 0.588 | 0.584 | 0.655 | 0.63174 | 0.58 | 0.62 | 0.62 | < 5 | 0.568 | 0.59 | 0.604 | [7 empty] | 0.62623
| 1 | 5 | 0.61 | 0.61 | 0.74 | 0.656 | 0.62 | 0.67183 | 0.618 | 0.676 | 0.622 | 0.603 | 0.645 | 0.62928 | 0.6 | 0.61 | 0.62 | < 5 | 0.566 | 0.68 | 0.604 | [7 empty] | 172
| 1 | 6 | 0.59 | 0.62 | 0.74 | 0.681 | 0.61 | 0.70667 | 0.628 | 0.686 | 0.611 | 0.604 | 0.647 | 0.6295 | 0.6 | 0.62 | 0.61 | < 5 | 0.557 | 0.69 | 0.635
| Mean |  | 0.60833 | 0.61167 | 0.73833 | 0.65517 | 0.60833 | 0.68199 | 0.63017 | 0.68117 | 0.602 | 0.609 | 0.6395 | 0.6494 | 0.59 | 0.615 | 0.62 |  | 0.55933 | 0.65833 | 0.62083
| Median |  | 0.61 | 0.61 | 0.74 | 0.6555 | 0.615 | 0.67919 | 0.6315 | 0.6825 | 0.5995 | 0.609 | 0.643 | 0.64635 | 0.595 | 0.615 | 0.62 |  | 0.5585 | 0.665 | 0.6225
| Std Dev. |  | 0.013292 | 0.01169 | 0.0040825 | 0.020663 | 0.019408 | 0.018826 | 0.0083046 | 0.01189 | 0.016456 | 0.015466 | 0.015123 | 0.021531 | 0.012649 | 0.010488 | 0.0063246 |  | 0.006802 | 0.035449 | 0.014552
| Rel.Std.Dev. |  | 2.185% | 1.911% | 0.5529% | 3.154% | 3.19% | 2.76% | 1.318% | 1.746% | 2.734% | 2.54% | 2.365% | 3.315% | 2.144% | 1.705% | 1.02% |  | 1.216% | 5.385% | 2.344%
| PDM3 |  | -2.858% | -2.326% | 17.9% | 4.621% | -2.858% | 8.904% | 0.6285% | 8.772% | -3.869% | -2.751% | 2.119% | 3.699% | -5.786% | -1.793% | -0.9949% |  | -10.68% | 5.126% | -0.8619%
| Z-Score (Absolute) |  | 0.45 | 0.33 | 4.21 | 1.23 | 0.45 | 2.19 | 0.33 | 2.16 | 0.67 | 0.42 | 0.67 | 1.02 | 1.1 | 0.21 | 0.03 | 67.28 | 2.2 | 1.34 | 0
| Analytical results for K in OREAS 550b (Certified Value 0.205 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.2099 | 0.19 | 0.24 | 0.278 | 0.258 | 0.21 | 0.19 | 0.20507 | 0.2 | 0.19 | 0.2 | 0.21 | 0.19 | 0.22339 | 0.18 | 0.2 | 0.4037 | 0.2 | 0.19 | 0.18 | 0.23564 | 0.23 | 0.475 | 0.177 |  |  | 1
| 1 | 2 | 0.213 | 0.21 | 0.26 | 0.289 | 0.271 | 0.21 | 0.21 | 0.2028 | 0.2 | 0.2 | 0.19 | 0.21 | 0.19 | 0.21523 | 0.18 | 0.2 | 0.4024 | 0.2 | 0.19 | 0.18 | 0.23251 | 0.2241 | 0.436 | 0.177 |  |  | #N/A
| 1 | 3 | 0.2023 | 0.22 | 0.25 | 0.27 | 0.231 | 0.21 | 0.18 | 0.20558 | 0.19 | 0.2 | 0.19 | 0.21 | 0.19 | 0.21589 | 0.17 | 0.21 | 0.3835 | 0.19 | 0.2 | 0.18 | 0.23091 | 0.2175 | 0.49 | 0.182 |  |  | 16
| 1 | 4 | 0.2022 | 0.22 | 0.24 | 0.266 | 0.221 | 0.21 | 0.2 | 0.2086 | 0.2 | 0.2 | 0.19 | 0.22 | 0.19 | 0.21796 | 0.18 | 0.2 | 0.4021 | 0.21 | 0.19 | 0.18 | 0.2385 | 0.264 | 0.44 | 0.176 |  |  | 0.20498
| 1 | 5 | 0.2133 | 0.19 | 0.24 | 0.28 | 0.251 | 0.21 | 0.21 | 0.2092 | 0.2 | 0.19 | 0.19 | 0.21 | 0.19 | 0.21408 | 0.17 | 0.2 | 0.4065 | 0.2 | 0.19 | 0.18 | 0.23769 | 0.2358 | 0.448 | 0.175 |  |  | 173
| 1 | 6 | 0.2093 | 0.19 | 0.26 | 0.287 | 0.209 | 0.2 | 0.18 | 0.2011 | 0.2 | 0.2 | 0.19 | 0.21 | 0.19 | 0.21157 | 0.17 | 0.2 | 0.404 | 0.21 | 0.19 | 0.18 | 0.23186 | 0.2391 | 0.472 | 0.179
| Mean |  | 0.20833 | 0.20333 | 0.24833 | 0.27833 | 0.24017 | 0.20833 | 0.195 | 0.20539 | 0.19833 | 0.19667 | 0.19167 | 0.21167 | 0.19 | 0.21635 | 0.175 | 0.20167 | 0.40037 | 0.20167 | 0.19167 | 0.18 | 0.23452 | 0.23508 | 0.46017 | 0.17767
| Median |  | 0.2096 | 0.2 | 0.245 | 0.279 | 0.241 | 0.21 | 0.195 | 0.20532 | 0.2 | 0.2 | 0.19 | 0.21 | 0.19 | 0.21556 | 0.175 | 0.2 | 0.40305 | 0.2 | 0.19 | 0.18 | 0.23407 | 0.2329 | 0.46 | 0.177
| Std Dev. |  | 0.004977 | 0.015055 | 0.0098319 | 0.0090921 | 0.023702 | 0.0040825 | 0.013784 | 0.003165 | 0.0040825 | 0.005164 | 0.0040825 | 0.0040825 | 3.0405e-17 | 0.0040431 | 0.0054772 | 0.0040825 | 0.008409 | 0.0075277 | 0.0040825 | 3.0405e-17 | 0.0032039 | 0.016177 | 0.021858 | 0.0025033
| Rel.Std.Dev. |  | 2.389% | 7.404% | 3.959% | 3.267% | 9.869% | 1.96% | 7.069% | 1.541% | 2.058% | 2.626% | 2.13% | 1.929% | 1.6e-14% | 1.869% | 3.13% | 2.024% | 2.1% | 3.733% | 2.13% | 1.689e-14% | 1.366% | 6.881% | 4.75% | 1.409%
| PDM3 |  | 1.634% | -0.8053% | 21.15% | 35.78% | 17.16% | 1.634% | -4.871% | 0.1975% | -3.245% | -4.058% | -6.497% | 3.26% | -7.31% | 5.546% | -14.63% | -1.618% | 95.32% | -1.618% | -6.497% | -12.19% | 14.41% | 14.68% | 124.5% | -13.33%
| Z-Score (Absolute) |  | 0.21 | 0.05 | 2.34 | 3.93 | 1.9 | 0.21 | 0.5 | 0.05 | 0.32 | 0.41 | 0.67 | 0.39 | 0.76 | 0.64 | 1.56 | 0.14 | 10.41 | 0.14 | 0.67 | 1.29 | 1.6 | 1.63 | 13.59 | 1.42
| Analytical results for La in OREAS 550b (Certified Value 19.2 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 1
| 1 | 1 | 19.981 | 17.4 | 22 | 27.5 | 43.0784 | 25.4 | 19 | 15.7 | 21.0266 | 19.7 | 18.9 | 18.6 | 21.4 | 19.2 | 21.2734 | 16.4 | 19.2 | 31.23 | 17 | 18 | 15.9 | 20.276 | 19 | 36.8 | 16.9 |  | 1
| 1 | 2 | 19.889 | 18.6 | 22.3 | 27 | 43.06 | 27.4 | 20 | 15.5 | 20.8268 | 20.1 | 19.9 | 19 | 21.5 | 19.4 | 20.7316 | 16.6 | 19.9 | 29.89 | 18 | 18 | 16.3 | 19.994 | < 1 | 34.3 | 16.6 |  | 7
| 1 | 3 | 20.127 | 19.5 | 21.4 | 27.1 | 43.0464 | 22.7 | 20 | 15 | 21.4716 | 19.3 | 19.6 | 19.2 | 21.3 | 18.9 | 21.9924 | 17.2 | 20.1 | 29.23 | 17 | 18 | 16.4 | 19.894 | 29 | 37.1 | 16.8 |  | 16
| 1 | 4 | 20.22 | 19.5 | 22.5 | 27.3 | 43.0256 | 23.6 | 20 | 15.7 | 21.7759 | 20 | 19.5 | 18.5 | 20.2 | 19.5 | 20.586 | 17.1 | 19.2 | 30.11 | 18 | 18 | 15.9 | 20.418 | < 1 | 38.4 | 16.3 |  | 19.2317
| 1 | 5 | 19.87 | 17 | 20.8 | 27.7 | 43.0928 | 25.1 | 19 | 15 | 21.5476 | 19.4 | 18.7 | 19.4 | 20.6 | 19.6 | 20.8744 | 17.4 | 19.8 | 28.8 | 19 | 18 | 15.7 | 20.482 | < 1 | 30.2 | 16.4 |  | 174
| 1 | 6 | 19.866 | 17.3 | 21.2 | 27.6 | 43.0536 | 21.2 | 20 | 15.2 | 21.3374 | 19.4 | 19.2 | 19.2 | 21.1 | 19.5 | 20.9922 | 17 | 19.8 | 28.95 | 19 | 18 | 16.2 | 20.136 | 8 | 39 | 17.5
| Mean |  | 19.9922 | 18.2167 | 21.7 | 27.3667 | 43.0595 | 24.2333 | 19.6667 | 15.35 | 21.331 | 19.65 | 19.3 | 18.9833 | 21.0167 | 19.35 | 21.075 | 16.95 | 19.6667 | 29.7017 | 18 | 18 | 16.0667 | 20.2 | 18.6667 | 35.9667 | 16.75
| Median |  | 19.935 | 18 | 21.7 | 27.4 | 43.0568 | 24.35 | 20 | 15.35 | 21.4045 | 19.55 | 19.35 | 19.1 | 21.2 | 19.45 | 20.9333 | 17.05 | 19.8 | 29.56 | 18 | 18 | 16.05 | 20.206 | 19 | 36.95 | 16.7
| Std Dev. |  | 0.14946 | 1.1339 | 0.66933 | 0.28048 | 0.023757 | 2.1951 | 0.5164 | 0.32711 | 0.34973 | 0.33912 | 0.45166 | 0.36009 | 0.51153 | 0.25884 | 0.50699 | 0.37815 | 0.37771 | 0.90978 | 0.89443 | 0 | 0.27325 | 0.23366 | 10.504 | 3.2599 | 0.43243
| Rel.Std.Dev. |  | 0.7476% | 6.224% | 3.084% | 1.025% | 0.05517% | 9.058% | 2.626% | 2.131% | 1.64% | 1.726% | 2.34% | 1.897% | 2.434% | 1.338% | 2.406% | 2.231% | 1.921% | 3.063% | 4.969% | 0% | 1.701% | 1.157% | 56.27% | 9.064% | 2.582%
| PDM3 |  | 3.954% | -5.278% | 12.83% | 42.3% | 123.9% | 26.01% | 2.262% | -20.18% | 10.92% | 2.175% | 0.3554% | -1.291% | 9.282% | 0.6153% | 9.585% | -11.86% | 2.262% | 54.44% | -6.404% | -6.404% | -16.46% | 5.035% | -2.938% | 87.02% | -12.9%
| Z-Score (Absolute) |  | 0.13 | 0.59 | 0.82 | 3.12 | 9.46 | 1.85 | 0 | 1.75 | 0.67 | 0.01 | 0.15 | 0.28 | 0.55 | 0.13 | 0.57 | 1.1 | 0 | 4.06 | 0.67 | 0.67 | 1.46 | 0.22 | 4.08 | 6.59 | 1.18
| Analytical results for Li in OREAS 550b (Certified Value 49.9 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 1
| 1 | 1 | 49.8 | 63.8 | 60.9 | 75.16 | 69 | 54 | 39 | 50.824 | 46 | 42.3 | 44 | 52.1 | 45.6 | 53.5389 | 41.3 | 52 | 53.3 | 29.8 | 52.66 | 50 | 57.9 | 43.3 |  |  |  |  | 1
| 1 | 2 | 49.7 | 49.3 | 58.6 | 75.41 | 67 | 54 | 43 | 49.1102 | 46.3 | 44.8 | 43.2 | 53.2 | 43.9 | 49.6181 | 40.3 | 53 | 50.7 | 29.5 | 52.36 | 47 | 54.3 | 41.1 |  |  |  |  | #N/A
| 1 | 3 | 50.1 | 50.2 | 59.8 | 74.71 | 61 | 55 | 40 | 50.6871 | 44.8 | 45.4 | 44.3 | 51.7 | 44.7 | 52.9007 | 42.1 | 53 | 47.6 | 29 | 54.23 | 45 | 55.5 | 43.7 |  |  |  |  | 16
| 1 | 4 | 51.1 | 59.8 | 59.9 | 73.5 | 60 | 55 | 42 | 50.9823 | 45.7 | 44.1 | 42 | 51.9 | 44.1 | 52.9423 | 42.1 | 53 | 46.3 | 29.2 | 56 | 47 | 55.7 | 41.6 |  |  |  |  | 49.8796
| 1 | 5 | 50.4 | 48.2 | 60.1 | 74.12 | 68 | 54 | 43 | 49.6445 | 45.3 | 42.8 | 42.5 | 51.4 | 44.5 | 51.5724 | 41.7 | 52 | 50.2 | 28.4 | 57.41 | 45 | 50.1 | 40.6 |  |  |  |  | 175
| 1 | 6 | 49.4 | 58.7 | 60.2 | 76.19 | 62 | 54 | 39 | 48.6897 | 45.9 | 44.4 | 42.4 | 50.7 | 43.8 | 51.186 | 40 | 53 | 47.2 | 29.3 | 55.3 | 52 | 55.1 | 39.4
| Mean |  | 50.0833 | 55 | 59.9167 | 74.8483 | 64.5 | 54.3333 | 41 | 49.9896 | 45.6667 | 43.9667 | 43.0667 | 51.8333 | 44.4333 | 51.9597 | 41.25 | 52.6667 | 49.2167 | 29.2 | 54.66 | 47.6667 | 54.7667 | 41.6167
| Median |  | 49.95 | 54.45 | 60 | 74.935 | 64.5 | 54 | 41 | 50.1658 | 45.8 | 44.25 | 42.85 | 51.8 | 44.3 | 52.2365 | 41.5 | 53 | 48.9 | 29.25 | 54.765 | 47 | 55.3 | 41.35
| Std Dev. |  | 0.6047 | 6.5718 | 0.75211 | 0.95711 | 3.937 | 0.5164 | 1.8974 | 0.97471 | 0.53914 | 1.1911 | 0.9288 | 0.82865 | 0.66833 | 1.4544 | 0.90719 | 0.5164 | 2.6468 | 0.47749 | 1.9618 | 2.8048 | 2.582 | 1.6364
| Rel.Std.Dev. |  | 1.207% | 11.95% | 1.255% | 1.279% | 6.104% | 0.9504% | 4.628% | 1.95% | 1.181% | 2.709% | 2.157% | 1.599% | 1.504% | 2.799% | 2.199% | 0.9805% | 5.378% | 1.635% | 3.589% | 5.884% | 4.715% | 3.932%
| PDM3 |  | 0.4084% | 10.27% | 20.12% | 50.06% | 29.31% | 8.929% | -17.8% | 0.2205% | -8.446% | -11.85% | -13.66% | 3.917% | -10.92% | 4.17% | -17.3% | 5.588% | -1.329% | -41.46% | 9.584% | -4.437% | 9.798% | -16.57%
| Z-Score (Absolute) |  | 0.01 | 0.69 | 1.37 | 3.45 | 2.01 | 0.6 | 1.26 | 0.01 | 0.61 | 0.84 | 0.97 | 0.25 | 0.78 | 0.27 | 1.22 | 0.37 | 0.11 | 2.9 | 0.64 | 0.33 | 0.66 | 1.17
| Analytical results for Lu in OREAS 550b (Certified Value 0.065 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 03 | 05 | 08 | 15 | 17 | 22 | 23 | 26 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | [17 empty] | 3
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.4g | [17 empty] | 3
| 1 | 1 | 0.067 | < 0.1 | 0.12808 | 0.07 | 0.073497 | 0.06 | < 5 | 0.061 | 0.06 | [17 empty] | 1
| 1 | 2 | 0.07 | < 0.1 | 0.12336 | 0.06 | 0.076331 | 0.06 | < 5 | 0.063 | 0.06 | [17 empty] | 8
| 1 | 3 | 0.066 | < 0.1 | 0.12792 | 0.06 | 0.074744 | 0.06 | < 5 | 0.062 | 0.07 | [17 empty] | 16
| 1 | 4 | 0.069 | < 0.1 | 0.13352 | 0.07 | 0.076684 | 0.06 | < 5 | 0.062 | 0.07 | [17 empty] | 0.065362
| 1 | 5 | 0.069 | < 0.1 | 0.12864 | 0.06 | 0.072064 | 0.06 | < 5 | 0.06 | 0.06 | [17 empty] | 176
| 1 | 6 | 0.068 | < 0.1 | 0.11472 | 0.06 | 0.073707 | 0.06 | < 5 | 0.059 | 0.07
| Mean |  | 0.068167 |  | 0.12604 | 0.063333 | 0.074505 | 0.06 |  | 0.061167 | 0.065
| Median |  | 0.0685 |  | 0.128 | 0.06 | 0.074226 | 0.06 |  | 0.0615 | 0.065
| Std Dev. |  | 0.001472 |  | 0.0064145 | 0.005164 | 0.001775 | 0 |  | 0.001472 | 0.0054772
| Rel.Std.Dev. |  | 2.159% |  | 5.089% | 8.154% | 2.382% | 0% |  | 2.406% | 8.427%
| PDM3 |  | 4.291% |  | 92.83% | -3.104% | 13.99% | -8.203% |  | -6.418% | -0.5537%
| Z-Score (Absolute) |  | 0.43 | 2.02 | 8.23 | 0.22 | 1.28 | 0.67 | 328.39 | 0.52 | 0
| Analytical results for Mg in OREAS 550b (Certified Value 3.53 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 3.42 | 3.39 | 3.76 | 3.48 | 4.067 | 3.54 | 3.21 | 3.4498 | 3.48 | 3.41 | 3.46 | 3.84 | 3.42 | 3.8118 | 3.43 | 3.65 | 3.74 | 3.3 | 3.46 | 2.91 | 3.5298 | 3.612 | 3.79 | 3.368 |  |  | 1
| 1 | 2 | 3.5 | 3.44 | 3.82 | 3.62 | 3.896 | 3.55 | 3.27 | 3.3497 | 3.51 | 3.52 | 3.42 | 3.89 | 3.31 | 3.6951 | 3.52 | 3.66 | 3.64 | 3.37 | 3.52 | 3.02 | 3.488 | 3.534 | 3.88 | 3.376 |  |  | #N/A
| 1 | 3 | 3.32 | 3.55 | 3.7 | 3.69 | 3.923 | 3.53 | 3.23 | 3.4265 | 3.47 | 3.52 | 3.42 | 3.79 | 3.34 | 3.7594 | 3.41 | 3.69 | 3.7 | 3.3 | 3.54 | 2.93 | 3.4852 | 3.468 | 3.98 | 3.426 |  |  | 16
| 1 | 4 | 3.33 | 3.58 | 3.54 | 3.58 | 3.83 | 3.52 | 3.27 | 3.4822 | 3.48 | 3.5 | 3.32 | 3.85 | 3.3 | 3.7851 | 3.52 | 3.69 | 3.75 | 3.45 | 3.53 | 2.9 | 3.6271 | 3.63 | 3.64 | 3.366 |  |  | 3.5321
| 1 | 5 | 3.49 | 3.31 | 3.75 | 3.69 | 4.121 | 3.53 | 3.25 | 3.4718 | 3.46 | 3.43 | 3.35 | 3.79 | 3.33 | 3.7101 | 3.5 | 3.64 | 3.77 | 3.43 | 3.5 | 2.99 | 3.6049 | 3.576 | 3.73 | 3.355 |  |  | 177
| 1 | 6 | 3.45 | 3.49 | 3.54 | 3.72 | 3.705 | 3.49 | 3.22 | 3.3232 | 3.46 | 3.51 | 3.34 | 3.77 | 3.3 | 3.7201 | 3.45 | 3.68 | 3.72 | 3.48 | 3.48 | 2.96 | 3.54 | 3.685 | 3.88 | 3.407
| Mean |  | 3.4183 | 3.46 | 3.685 | 3.63 | 3.9237 | 3.5267 | 3.2417 | 3.4172 | 3.4767 | 3.4817 | 3.385 | 3.8217 | 3.3333 | 3.7469 | 3.4717 | 3.6683 | 3.72 | 3.3883 | 3.505 | 2.9517 | 3.5458 | 3.5842 | 3.8167 | 3.383
| Median |  | 3.435 | 3.465 | 3.725 | 3.655 | 3.9095 | 3.53 | 3.24 | 3.4381 | 3.475 | 3.505 | 3.385 | 3.815 | 3.32 | 3.7398 | 3.475 | 3.67 | 3.73 | 3.4 | 3.51 | 2.945 | 3.5349 | 3.594 | 3.835 | 3.372
| Std Dev. |  | 0.077825 | 0.10119 | 0.11862 | 0.089889 | 0.15287 | 0.020656 | 0.025626 | 0.065949 | 0.018619 | 0.048751 | 0.055767 | 0.045789 | 0.045461 | 0.04598 | 0.047924 | 0.02137 | 0.046043 | 0.077309 | 0.030822 | 0.047081 | 0.058972 | 0.076316 | 0.12176 | 0.027466
| Rel.Std.Dev. |  | 2.277% | 2.925% | 3.219% | 2.476% | 3.896% | 0.5857% | 0.7905% | 1.93% | 0.5355% | 1.4% | 1.647% | 1.198% | 1.364% | 1.227% | 1.38% | 0.5825% | 1.238% | 2.282% | 0.8794% | 1.595% | 1.663% | 2.129% | 3.19% | 0.8119%
| PDM3 |  | -3.222% | -2.042% | 4.328% | 2.77% | 11.08% | -0.1551% | -8.224% | -3.255% | -1.571% | -1.429% | -4.166% | 8.197% | -5.629% | 6.081% | -1.712% | 3.856% | 5.318% | -4.071% | -0.7685% | -16.43% | 0.3878% | 1.473% | 8.055% | -4.222%
| Z-Score (Absolute) |  | 0.46 | 0.21 | 1.18 | 0.84 | 2.65 | 0.21 | 1.55 | 0.47 | 0.1 | 0.07 | 0.67 | 2.02 | 0.99 | 1.56 | 0.13 | 1.08 | 1.4 | 0.65 | 0.07 | 3.34 | 0.32 | 0.56 | 1.99 | 0.68
| Analytical results for Mn in OREAS 550b (Certified Value 0.018 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.0181 | 0.0181 | 0.0171 | 0.0168 | 0.0237 | 0.0173 | 0.0192 | 0.018732 | 0.0176 | 0.0185 | 0.0178 | 0.0182 | 0.018 | 0.01939 | 0.0166 | 0.0179 | 0.0195 | 0.0182 | 0.0174 | 0.01742 | 0.016757 | 0.0194 | 0.0162 | 0.0191 |  |  | 1
| 1 | 2 | 0.0184 | 0.018 | 0.0178 | 0.0172 | 0.0221 | 0.0179 | 0.0185 | 0.018778 | 0.018 | 0.0189 | 0.0177 | 0.0184 | 0.0175 | 0.018529 | 0.0171 | 0.0178 | 0.0189 | 0.018 | 0.0175 | 0.01759 | 0.01635 | 0.0194 | 0.0169 | 0.0178 |  |  | #N/A
| 1 | 3 | 0.0176 | 0.0183 | 0.0179 | 0.0174 | 0.0219 | 0.018 | 0.0191 | 0.019174 | 0.0179 | 0.0188 | 0.0175 | 0.0181 | 0.0176 | 0.019024 | 0.017 | 0.0179 | 0.019 | 0.0181 | 0.0174 | 0.01708 | 0.016335 | 0.0201 | 0.0171 | 0.0188 |  |  | 16
| 1 | 4 | 0.0172 | 0.0181 | 0.017 | 0.0169 | 0.0215 | 0.0175 | 0.0185 | 0.019054 | 0.0178 | 0.0189 | 0.0171 | 0.0184 | 0.0174 | 0.01883 | 0.017 | 0.0179 | 0.0196 | 0.0184 | 0.0172 | 0.01705 | 0.01703 | 0.0201 | 0.0172 | 0.018 |  |  | 0.017955
| 1 | 5 | 0.0181 | 0.0177 | 0.0175 | 0.0172 | 0.0201 | 0.0183 | 0.0192 | 0.018525 | 0.0177 | 0.0186 | 0.0172 | 0.018 | 0.0175 | 0.018486 | 0.0173 | 0.0177 | 0.0196 | 0.018 | 0.0175 | 0.0173 | 0.016878 | 0.0194 | 0.0167 | 0.0178 |  |  | 178
| 1 | 6 | 0.018 | 0.0183 | 0.0168 | 0.0174 | 0.0211 | 0.0182 | 0.0187 | 0.018659 | 0.0175 | 0.0189 | 0.0173 | 0.018 | 0.0174 | 0.018478 | 0.0172 | 0.0178 | 0.0193 | 0.0186 | 0.0176 | 0.01707 | 0.016511 | 0.0201 | 0.0167 | 0.0174
| Mean |  | 0.0179 | 0.018083 | 0.01735 | 0.01715 | 0.021733 | 0.017867 | 0.018867 | 0.01882 | 0.01775 | 0.018767 | 0.017433 | 0.018183 | 0.017567 | 0.018789 | 0.017033 | 0.017833 | 0.019317 | 0.018217 | 0.017433 | 0.017252 | 0.016644 | 0.01975 | 0.0168 | 0.01815
| Median |  | 0.01805 | 0.0181 | 0.0173 | 0.0172 | 0.0217 | 0.01795 | 0.0189 | 0.018755 | 0.01775 | 0.01885 | 0.0174 | 0.01815 | 0.0175 | 0.01868 | 0.01705 | 0.01785 | 0.0194 | 0.01815 | 0.01745 | 0.01719 | 0.016634 | 0.01975 | 0.0168 | 0.0179
| Std Dev. |  | 0.00042895 | 0.00022286 | 0.00045056 | 0.000251 | 0.0011961 | 0.00039328 | 0.00033862 | 0.00024591 | 0.00018708 | 0.00017512 | 0.00028048 | 0.00018348 | 0.00022509 | 0.00036711 | 0.00024221 | 8.165e-05 | 0.00030605 | 0.00024014 | 0.00013663 | 0.00022284 | 0.00028846 | 0.00038341 | 0.00035777 | 0.00065651
| Rel.Std.Dev. |  | 2.396% | 1.232% | 2.597% | 1.464% | 5.504% | 2.201% | 1.795% | 1.307% | 1.054% | 0.9331% | 1.609% | 1.009% | 1.281% | 1.954% | 1.422% | 0.4578% | 1.584% | 1.318% | 0.7837% | 1.292% | 1.733% | 1.941% | 2.13% | 3.617%
| PDM3 |  | -0.3036% | 0.7175% | -3.367% | -4.481% | 21.05% | -0.4893% | 5.08% | 4.822% | -1.139% | 4.523% | -2.903% | 1.274% | -2.16% | 4.651% | -5.131% | -0.6749% | 7.587% | 1.46% | -2.903% | -3.915% | -7.301% | 10% | -6.43% | 1.089%
| Z-Score (Absolute) |  | 0.02 | 0.23 | 0.62 | 0.85 | 4.46 | 0.02 | 1.14 | 1.08 | 0.15 | 1.02 | 0.52 | 0.35 | 0.37 | 1.05 | 0.98 | 0.06 | 1.66 | 0.39 | 0.52 | 0.73 | 1.43 | 2.16 | 1.25 | 0.31
| Analytical results for Mo in OREAS 550b (Certified Value 6.56 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 7 | 6.27 | 6.33 | 6.6 | 7.14 | 7.07 | 6.2554 | 7.04 | 6.49 | 6.73 | 6.04 | 6.33 | 6.9705 | 6.53 | 6.38 | 6.73 | 6.8 | 6.1 | 6.5 | 5.6 | 6.582 | 5.56 | 6.3 | 6.67 |  |  | 1
| 1 | 2 | 6.8 | 6.27 | 6.77 | 6.4 | 6.87 | 7.12 | 6.1822 | 7.15 | 6.48 | 6.89 | 6.17 | 6.41 | 6.7592 | 6.63 | 6.67 | 6.54 | 6.8 | 6.1 | 6.5 | 5.8 | 6.434 | 5.47 | 6.4 | 6.32 |  |  | 9
| 1 | 3 | 6.9 | 6.33 | 6.7 | 6.6 | 6.31 | 6.89 | 6.4169 | 6.94 | 6.59 | 7.11 | 6.37 | 6.29 | 7.0074 | 6.41 | 6.52 | 6.53 | 6.8 | 6 | 6.5 | 5.8 | 6.427 | 5.66 | 6.3 | 6.8 |  |  | 16
| 1 | 4 | 6.9 | 6.42 | 6.44 | 6.5 | 6.45 | 7.06 | 6.3838 | 7.13 | 6.55 | 6.9 | 5.89 | 6.54 | 7.0697 | 6.31 | 6.42 | 6.66 | 6.6 | 6.2 | 6.4 | 5.7 | 6.785 | 5.72 | 5.9 | 6.53 |  |  | 6.5632
| 1 | 5 | 6.8 | 6.3 | 6.55 | 6.6 | 6.5 | 6.95 | 6.211 | 6.9 | 6.45 | 7.12 | 6.26 | 6.46 | 6.9748 | 6.43 | 6.51 | 6.5 | 6.7 | 6.2 | 6.4 | 5.5 | 6.642 | 5.92 | 6.4 | 6.39 |  |  | 179
| 1 | 6 | 6.7 | 6.26 | 6.26 | 6.5 | 6.38 | 7.02 | 6.3417 | 6.83 | 6.6 | 6.98 | 5.99 | 6.47 | 6.9805 | 6.18 | 6.51 | 6.5 | 6.7 | 6.3 | 6.5 | 5.5 | 6.461 | 5.66 | 6.5 | 6.35
| Mean |  | 6.85 | 6.3083 | 6.5083 | 6.5333 | 6.6083 | 7.0183 | 6.2985 | 6.9983 | 6.5267 | 6.955 | 6.12 | 6.4167 | 6.9603 | 6.415 | 6.5017 | 6.5767 | 6.7333 | 6.15 | 6.4667 | 5.65 | 6.5552 | 5.665 | 6.3 | 6.51
| Median |  | 6.85 | 6.285 | 6.495 | 6.55 | 6.475 | 7.04 | 6.2985 | 6.99 | 6.52 | 6.94 | 6.105 | 6.435 | 6.9776 | 6.42 | 6.51 | 6.535 | 6.75 | 6.15 | 6.5 | 5.65 | 6.5215 | 5.66 | 6.35 | 6.46
| Std Dev. |  | 0.10488 | 0.06047 | 0.20253 | 0.08165 | 0.32529 | 0.084715 | 0.09613 | 0.12922 | 0.062183 | 0.14816 | 0.17933 | 0.09331 | 0.10521 | 0.15871 | 0.10028 | 0.095638 | 0.08165 | 0.10488 | 0.05164 | 0.13784 | 0.14217 | 0.15307 | 0.20976 | 0.19298
| Rel.Std.Dev. |  | 1.531% | 0.9586% | 3.112% | 1.25% | 4.922% | 1.207% | 1.526% | 1.846% | 0.9527% | 2.13% | 2.93% | 1.454% | 1.511% | 2.474% | 1.542% | 1.454% | 1.213% | 1.705% | 0.7986% | 2.44% | 2.169% | 2.702% | 3.33% | 2.964%
| PDM3 |  | 4.37% | -3.883% | -0.8361% | -0.4552% | 0.6875% | 6.934% | -4.034% | 6.63% | -0.5568% | 5.969% | -6.753% | -2.233% | 6.051% | -2.258% | -0.9377% | 0.205% | 2.592% | -6.296% | -1.471% | -13.91% | -0.1226% | -13.69% | -4.01% | -0.8107%
| Z-Score (Absolute) |  | 1.12 | 0.66 | 0 | 0.08 | 0.33 | 1.67 | 0.69 | 1.61 | 0.06 | 1.47 | 1.28 | 0.3 | 1.48 | 0.31 | 0.02 | 0.22 | 0.74 | 1.18 | 0.14 | 2.83 | 0.15 | 2.78 | 0.69 | 0
| Analytical results for Na in OREAS 550b (Certified Value 0.01 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 3
| 1 | 1 | 0.006 | 0.018 | 0.01 | 0.017 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.02 | < 0.01 | 0.02 | 0.01 | 0.006 | 0.01 | 0.01 | 0.012257 | 0.015 | 0.0077 | 0.012 |  |  |  |  | 1
| 1 | 2 | 0.006 | 0.012 | 0.01 | 0.017 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | 0.02 | 0.01 | 0.006 | 0.01 | 0.01 | 0.0099348 | 0.015 | 0.008 | 0.011 |  |  |  |  | 21
| 1 | 3 | 0.006 | 0.013 | 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.02 | 0.01 | 0.008 | < 0.01 | 0.01 | 0.013347 | 0.015 | 0.0085 | 0.013 |  |  |  |  | 16
| 1 | 4 | 0.006 | 0.017 | < 0.01 | 0.016 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.02 | 0.01 | 0.007 | 0.02 | 0.01 | 0.0098374 | 0.03 | 0.0095 | 0.011 |  |  |  |  | 0.0095787
| 1 | 5 | 0.006 | 0.012 | 0.01 | 0.016 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.02 | 0.01 | 0.008 | 0.02 | 0.01 | 0.0097829 | 0.007 | 0.009 | 0.01 |  |  |  |  | 180
| 1 | 6 | 0.011 | 0.018 | 0.01 | 0.015 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.02 | 0.01 | 0.008 | 0.01 | 0.01 | 0.0093982 | 0.007 | 0.0094 | 0.01
| Mean |  | 0.0068333 | 0.015 | 0.01 | 0.016833 | 0.01 | 0.01 |  | 0.01 | 0.01 | 0.013333 | 0.01 | 0.011667 |  | 0.02 | 0.01 | 0.0071667 | 0.014 | 0.01 | 0.010759 | 0.014833 | 0.0086833 | 0.011167
| Median |  | 0.006 | 0.015 | 0.01 | 0.0165 | 0.01 | 0.01 |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 |  | 0.02 | 0.01 | 0.0075 | 0.01 | 0.01 | 0.0098861 | 0.015 | 0.00875 | 0.011
| Std Dev. |  | 0.0020412 | 0.0029665 | 0 | 0.0017224 | 0 | 0 |  | 0 | 0 | 0.005164 | 0 | 0.0040825 |  | 0 | 0 | 0.00098319 | 0.0054772 | 0 | 0.0016293 | 0.0084004 | 0.0007414 | 0.001169
| Rel.Std.Dev. |  | 29.87% | 19.78% | 0% | 10.23% | 0% | 0% |  | 0% | 0% | 38.73% | 0% | 34.99% |  | 0% | 0% | 13.72% | 39.12% | 0% | 15.14% | 56.63% | 8.538% | 10.47%
| PDM3 |  | -28.66% | 56.6% | 4.399% | 75.74% | 4.399% | 4.399% |  | 4.399% | 4.399% | 39.2% | 4.399% | 21.8% |  | 108.8% | 4.399% | -25.18% | 46.16% | 4.399% | 12.33% | 54.86% | -9.347% | 16.58%
| Z-Score (Absolute) |  | 1.43 | 2.26 | 0.38 | 3.09 | 0 | 0 | 2.26 | 0.38 | 0 | 1.51 | 0 | 0.75 | 2.26 | 4.52 | 0 | 1.28 | 1.13 | 0 | 0.34 | 2.19 | 0.6 | 0.53
| Analytical results for Nb in OREAS 550b (Indicative Value 0.048 ppm) | [27 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 21 | 22 | 24 | 25 | 26 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [8 empty] | 3
| 1 | 1 | 0.03 | < 0.1 | < 0.2 | 0.05 | 0.07 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.022148 | < 0.05 | < 0.1 | < 5 | 0.04 | 0.07 | 0.08 | [8 empty] | 1
| 1 | 2 | 0.03 | < 0.1 | < 0.2 | < 0.05 | 0.05 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.015118 | < 0.05 | < 0.1 | < 5 | 0.03 | 0.05 | 0.07 | [8 empty] | 25
| 1 | 3 | 0.03 | < 0.1 | < 0.2 | < 0.05 | 0.08 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.01473 | < 0.05 | < 0.1 | < 5 | 0.04 | 0.06 | 0.07 | [8 empty] | 16
| 1 | 4 | 0.03 | < 0.1 | < 0.2 | 0.06 | 0.06 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.018601 | < 0.05 | < 0.1 | < 5 | 0.03 | 0.06 | 0.07 | [8 empty] | 0.048032
| 1 | 5 | 0.03 | < 0.1 | < 0.2 | < 0.05 | 0.05 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.021447 | < 0.05 | < 0.1 | < 5 | 0.05 | 0.06 | 0.07 | [8 empty] | 31
| 1 | 6 | 0.03 | < 0.1 | < 0.2 | 0.05 | 0.08 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.025308 | < 0.05 | < 0.1 | < 5 | 0.03 | 0.06 | 0.07
| Mean |  | 0.03 |  |  | 0.053333 | 0.065 | [6 empty] | 0.019558 |  |  |  | 0.036667 | 0.06 | 0.071667
| Median |  | 0.03 |  |  | 0.05 | 0.065 | [6 empty] | 0.020024 |  |  |  | 0.035 | 0.06 | 0.07
| Std Dev. |  | 0 |  |  | 0.0057735 | 0.013784 | [6 empty] | 0.004178 |  |  |  | 0.008165 | 0.0063246 | 0.0040825
| Rel.Std.Dev. |  | 0% |  |  | 10.83% | 21.21% | [6 empty] | 21.36% |  |  |  | 22.27% | 10.54% | 5.696%
| PDM3 |  | -37.54% |  |  | 11.04% | 35.33% | [6 empty] | -59.28% |  |  |  | -23.66% | 24.92% | 49.21%
| Z-Score (Absolute) |  | 0.41 | 0.63 | 3.24 | 0.07 | 1.41 | 3.24 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.96 | 0.67 | 0.63 | 128.53 | 0.07 | 1.15 | 1.76
| Analytical results for Nd in OREAS 550b (Certified Value 17.5 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 1
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 1
| 1 | 1 | 18.724 | 21 | 37.2656 | 21.3462 | 18.187 | 12.1 | 14.4 | [19 empty] | 1
| 1 | 2 | 18.771 | 20.4 | 37.304 | 20.5088 | 18.203 | 11.9 | 14.54 | [19 empty] | 12
| 1 | 3 | 18.621 | 20.4 | 37.2896 | 20.8624 | 17.963 | 12.4 | 15.03 | [19 empty] | 16
| 1 | 4 | 19.008 | 21.2 | 37.3088 | 20.8 | 18.815 | 12.6 | 14.46 | [19 empty] | 17.4893
| 1 | 5 | 18.945 | 19.7 | 37.3352 | 19.7281 | 18.452 | 12.8 | 14.01 | [19 empty] | 181
| 1 | 6 | 18.596 | 20.9 | 37.3288 | 19.6841 | 18.07 | 12.3 | 14.19
| Mean |  | 18.7775 | 20.6 | 37.3053 | 20.4883 | 18.2817 | 12.35 | 14.4383
| Median |  | 18.7475 | 20.65 | 37.3064 | 20.6544 | 18.195 | 12.35 | 14.43
| Std Dev. |  | 0.16824 | 0.54772 | 0.025623 | 0.66306 | 0.30821 | 0.32711 | 0.34856
| Rel.Std.Dev. |  | 0.896% | 2.659% | 0.06869% | 3.236% | 1.686% | 2.649% | 2.414%
| PDM3 |  | 7.366% | 17.79% | 113.3% | 17.15% | 4.531% | -29.39% | -17.44%
| Z-Score (Absolute) |  | 0 | 0.67 | 6.86 | 0.63 | 0.18 | 2.38 | 1.61
| Analytical results for Ni in OREAS 550b (Certified Value 60 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 1
| 1 | 1 | 65.2 | 55.3 | 62.5 | 62 | 56.57 | 83 | 57 | 65.4 | 60.3274 | 63.3 | 57.6 | 61.3 | 61.3 | 57.8 | 63.0476 | 57.3 | 59.8 | 68 | 58.8 | 57.5 | 48.5 | 58.829 | 50.9 | 60.9 | 60 |  | 1
| 1 | 2 | 63.9 | 56.2 | 60.3 | 61 | 58 | 87 | 57 | 67.4 | 59.8096 | 61.8 | 59.1 | 60.7 | 62.5 | 57.6 | 59.9736 | 57.6 | 63.2 | 67 | 59.3 | 59.2 | 48.48 | 58.742 | 51.4 | 62.6 | 62 |  | 28
| 1 | 3 | 64.8 | 58.7 | 60.7 | 61 | 58.85 | 85 | 56 | 66.7 | 60.653 | 58.9 | 58.8 | 63.4 | 62.8 | 58.4 | 60.8307 | 58.3 | 62 | 67 | 58.9 | 58.4 | 47.05 | 58.734 | 51.1 | 59.1 | 63 |  | 16
| 1 | 4 | 65.6 | 57.6 | 60.5 | 62 | 59.68 | 81 | 56 | 65.5 | 61.3917 | 59.4 | 58.5 | 60.6 | 59.3 | 60 | 61.0273 | 56.7 | 61.6 | 69 | 60.7 | 58.2 | 46.57 | 61.709 | 52.6 | 56.5 | 61 |  | 60.0636
| 1 | 5 | 65.5 | 54.1 | 61 | 61 | 59.51 | 86 | 56 | 66.9 | 59.2356 | 58.8 | 57.1 | 62.5 | 61.3 | 61.4 | 60.1621 | 58.5 | 62.9 | 68 | 62 | 59.4 | 47.65 | 60.429 | 54.3 | 61.2 | 61 |  | 182
| 1 | 6 | 63 | 58.2 | 61.7 | 62 | 61.13 | 78 | 56 | 65.6 | 58.5947 | 59.2 | 57.6 | 62.7 | 62.7 | 58.8 | 59.6202 | 57.8 | 63 | 67 | 63.3 | 58.5 | 46.36 | 60.171 | 52.7 | 62.4 | 61
| Mean |  | 64.6667 | 56.6833 | 61.1167 | 61.5 | 58.9567 | 83.3333 | 56.3333 | 66.25 | 60.002 | 60.2333 | 58.1167 | 61.8667 | 61.65 | 59 | 60.7769 | 57.7 | 62.0833 | 67.6667 | 60.5 | 58.5333 | 47.435 | 59.769 | 52.1667 | 60.45 | 61.3333
| Median |  | 65 | 56.9 | 60.85 | 61.5 | 59.18 | 84 | 56 | 66.15 | 60.0685 | 59.3 | 58.05 | 61.9 | 61.9 | 58.6 | 60.4964 | 57.7 | 62.45 | 67.5 | 60 | 58.45 | 47.35 | 59.5 | 52 | 61.05 | 61
| Std Dev. |  | 1.0231 | 1.7904 | 0.83526 | 0.54772 | 1.5605 | 3.3862 | 0.5164 | 0.85499 | 1.0066 | 1.8683 | 0.79352 | 1.1605 | 1.3353 | 1.4533 | 1.2315 | 0.6603 | 1.2813 | 0.8165 | 1.845 | 0.69186 | 0.93001 | 1.2141 | 1.2925 | 2.3072 | 1.0328
| Rel.Std.Dev. |  | 1.582% | 3.159% | 1.367% | 0.8906% | 2.647% | 4.063% | 0.9167% | 1.291% | 1.678% | 3.102% | 1.365% | 1.876% | 2.166% | 2.463% | 2.026% | 1.144% | 2.064% | 1.207% | 3.05% | 1.182% | 1.961% | 2.031% | 2.478% | 3.817% | 1.684%
| PDM3 |  | 7.664% | -5.628% | 1.753% | 2.392% | -1.843% | 38.74% | -6.21% | 10.3% | -0.1025% | 0.2827% | -3.241% | 3.002% | 2.641% | -1.771% | 1.188% | -3.935% | 3.363% | 12.66% | 0.7266% | -2.548% | -21.03% | -0.4904% | -13.15% | 0.6434% | 2.114%
| Z-Score (Absolute) |  | 1.9 | 1.7 | 0.3 | 0.47 | 0.67 | 10.33 | 1.86 | 2.62 | 0.2 | 0.1 | 1.05 | 0.64 | 0.54 | 0.65 | 0.15 | 1.24 | 0.74 | 3.26 | 0.02 | 0.87 | 5.88 | 0.31 | 3.74 | 0 | 0.4
| Analytical results for P in OREAS 550b (Certified Value 0.039 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.04 | 0.036 | 0.045 | 0.0389 | 0.042 | 0.0373 | 0.0364 | 0.039341 | 0.039 | 0.039 | 0.038 | 0.042 | 0.039 | 0.040748 | 0.0363 | 0.04 | 0.0399 | 0.035 | 0.041 | 0.04 | 0.037923 | < 0.0009 | 0.0444 | 0.041 |  |  | 1
| 1 | 2 | 0.0394 | 0.037 | 0.044 | 0.0376 | 0.053 | 0.0388 | 0.038 | 0.038376 | 0.04 | 0.039 | 0.038 | 0.043 | 0.038 | 0.039504 | 0.0368 | 0.04 | 0.0393 | 0.036 | 0.042 | 0.04 | 0.036919 | < 0.0009 | 0.0429 | 0.04 |  |  | #N/A
| 1 | 3 | 0.0381 | 0.037 | 0.043 | 0.0389 | 0.037 | 0.0392 | 0.036 | 0.039811 | 0.039 | 0.04 | 0.037 | 0.042 | 0.038 | 0.041985 | 0.0368 | 0.04 | 0.04 | 0.036 | 0.042 | 0.04 | 0.035357 | < 0.0009 | 0.0479 | 0.041 |  |  | 16
| 1 | 4 | 0.0383 | 0.037 | 0.046 | 0.0388 | 0.038 | 0.0382 | 0.0382 | 0.040489 | 0.039 | 0.04 | 0.036 | 0.043 | 0.038 | 0.041623 | 0.0373 | 0.04 | 0.0415 | 0.036 | 0.04 | 0.04 | 0.037185 | < 0.0009 | 0.0451 | 0.039 |  |  | 0.039341
| 1 | 5 | 0.0394 | 0.035 | 0.043 | 0.0394 | 0.04 | 0.0373 | 0.0373 | 0.040421 | 0.039 | 0.038 | 0.037 | 0.041 | 0.038 | 0.040237 | 0.0368 | 0.04 | 0.0432 | 0.036 | 0.042 | 0.04 | 0.036233 | < 0.0009 | 0.0443 | 0.039 |  |  | 183
| 1 | 6 | 0.0389 | 0.036 | 0.045 | 0.0398 | 0.039 | 0.0371 | 0.0372 | 0.038032 | 0.039 | 0.039 | 0.037 | 0.042 | 0.038 | 0.040777 | 0.0371 | 0.04 | 0.0399 | 0.038 | 0.041 | 0.04 | 0.035997 | < 0.0009 | 0.0446 | 0.04
| Mean |  | 0.039017 | 0.036333 | 0.044333 | 0.0389 | 0.0415 | 0.037983 | 0.037183 | 0.039412 | 0.039167 | 0.039167 | 0.037167 | 0.042167 | 0.038167 | 0.040812 | 0.03685 | 0.04 | 0.040633 | 0.036167 | 0.041333 | 0.04 | 0.036602 |  | 0.044867 | 0.04
| Median |  | 0.03915 | 0.0365 | 0.0445 | 0.0389 | 0.0395 | 0.03775 | 0.03725 | 0.039576 | 0.039 | 0.039 | 0.037 | 0.042 | 0.038 | 0.040762 | 0.0368 | 0.04 | 0.03995 | 0.036 | 0.0415 | 0.04 | 0.036576 |  | 0.0445 | 0.04
| Std Dev. |  | 0.00072503 | 0.0008165 | 0.0012111 | 0.00074297 | 0.0058907 | 0.00088412 | 0.00086352 | 0.0010317 | 0.00040825 | 0.00075277 | 0.00075277 | 0.00075277 | 0.00040825 | 0.00090349 | 0.00033912 | 0 | 0.0014556 | 0.00098319 | 0.0008165 | 0 | 0.00092055 |  | 0.0016573 | 0.00089443
| Rel.Std.Dev. |  | 1.858% | 2.247% | 2.732% | 1.91% | 14.19% | 2.328% | 2.322% | 2.618% | 1.042% | 1.922% | 2.025% | 1.785% | 1.07% | 2.214% | 0.9203% | 0% | 3.582% | 2.719% | 1.975% | 0% | 2.515% |  | 3.694% | 2.236%
| PDM3 |  | -0.8255% | -7.646% | 12.69% | -1.122% | 5.487% | -3.452% | -5.486% | 0.1785% | -0.4442% | -0.4442% | -5.528% | 7.181% | -2.986% | 3.739% | -6.333% | 1.674% | 3.284% | -8.07% | 5.063% | 1.674% | -6.962% |  | 14.04% | 1.674%
| Z-Score (Absolute) |  | 0.06 | 1.05 | 1.92 | 0.1 | 0.87 | 0.44 | 0.74 | 0.09 | 0 | 0 | 0.74 | 1.11 | 0.37 | 0.61 | 0.86 | 0.31 | 0.55 | 1.11 | 0.81 | 0.31 | 0.95 | 14.39 | 2.12 | 0.31
| Analytical results for Pb in OREAS 550b (Certified Value 9.13 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 2
| 1 | 1 | 9.3 | 8.14 | 10.5 | 10.7 | 8 | 9 | 7.5 | 9.4959 | 9 | 8 | 8.3 | 9.8 | 8.7 | 10.546 | 8.2 | 8.8 | 11.9 | 14 | 8 | 8.2 | 14.3 | 11.66 | 6.89 | 11.6 | 8 |  | 1
| 1 | 2 | 9.1 | 8.64 | 9.8 | 10.6 | 8 | 9 | 7.8 | 9.2385 | 8.9 | 8.5 | 8.4 | 9.6 | 8.3 | 9.9319 | 8.2 | 8.9 | 11.2 | 14 | 8.4 | 8.1 | 14.2 | 11.97 | 6.79 | 11.7 | 8 |  | 13
| 1 | 3 | 9.3 | 8.74 | 9.7 | 10.6 | 6 | 9 | 7.4 | 9.5524 | 9.5 | 8.5 | 8.7 | 9.6 | 8.2 | 10.241 | 8.3 | 9 | 11.3 | 12 | 8.1 | 8.6 | 13.7 | 11.34 | 7.02 | 11.9 | 8 |  | 16
| 1 | 4 | 9.3 | 9.15 | 10.3 | 10.6 | 7 | 9 | 7.7 | 9.5224 | 9 | 8.4 | 8.2 | 8.9 | 8.6 | 10.2581 | 8.2 | 8.9 | 11.4 | 10 | 8.4 | 8.1 | 13.9 | 12.44 | 7.2 | 11 | 8 |  | 9.1344
| 1 | 5 | 9.2 | 7.72 | 9.5 | 10.4 | 8 | 9 | 7.4 | 10.9811 | 8.8 | 8.1 | 8.5 | 9.3 | 8.5 | 10.0811 | 8.1 | 8.7 | 11 | 11 | 8.7 | 8 | 13.8 | 10.42 | 7.28 | 11 | 8 |  | 184
| 1 | 6 | 9 | 7.59 | 10.2 | 10.6 | 5 | 9 | 7.6 | 11.0159 | 8.8 | 8.4 | 8.6 | 9.6 | 8.6 | 10.1133 | 8.1 | 9.1 | 11.1 | 10 | 8.7 | 8 | 13.9 | 10.83 | 7.05 | 11.7 | 8
| Mean |  | 9.2 | 8.33 | 10 | 10.5833 | 7 | 9 | 7.5667 | 9.9677 | 9 | 8.3167 | 8.45 | 9.4667 | 8.4833 | 10.1952 | 8.1833 | 8.9 | 11.3167 | 11.8333 | 8.3833 | 8.1667 | 13.9667 | 11.4433 | 7.0383 | 11.4833 | 8
| Median |  | 9.25 | 8.39 | 10 | 10.6 | 7.5 | 9 | 7.55 | 9.5374 | 8.95 | 8.4 | 8.45 | 9.6 | 8.55 | 10.1772 | 8.2 | 8.9 | 11.25 | 11.5 | 8.4 | 8.1 | 13.9 | 11.5 | 7.035 | 11.65 | 8
| Std Dev. |  | 0.12649 | 0.61521 | 0.38987 | 0.098319 | 1.2649 | 0 | 0.1633 | 0.80629 | 0.26077 | 0.2137 | 0.18708 | 0.32042 | 0.19408 | 0.20887 | 0.075277 | 0.14142 | 0.31885 | 1.8348 | 0.29269 | 0.22509 | 0.23381 | 0.74185 | 0.18368 | 0.38687 | 0
| Rel.Std.Dev. |  | 1.375% | 7.385% | 3.899% | 0.929% | 18.07% | 0% | 2.158% | 8.089% | 2.897% | 2.57% | 2.214% | 3.385% | 2.288% | 2.049% | 0.9199% | 1.589% | 2.818% | 15.51% | 3.491% | 2.756% | 1.674% | 6.483% | 2.61% | 3.369% | 0%
| PDM3 |  | 0.7186% | -8.806% | 9.477% | 15.86% | -23.37% | -1.471% | -17.16% | 9.123% | -1.471% | -8.952% | -7.492% | 3.638% | -7.127% | 11.61% | -10.41% | -2.566% | 23.89% | 29.55% | -8.222% | -10.59% | 52.9% | 25.28% | -22.95% | 25.72% | -12.42%
| Z-Score (Absolute) |  | 0.14 | 0.47 | 0.7 | 1.1 | 1.39 | 0 | 1 | 0.67 | 0 | 0.48 | 0.38 | 0.33 | 0.36 | 0.83 | 0.57 | 0.07 | 1.61 | 1.97 | 0.43 | 0.58 | 3.46 | 1.7 | 1.37 | 1.73 | 0.7
| Analytical results for Pd in OREAS 550b (Indicative Value < 10 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | [25 empty] | 1
| No | No | 09 | [25 empty] | ppb
|  |  | AR*OES/MS | [25 empty] | 1
|  |  | 01g | [25 empty] | 1
| 1 | 1 | < 10 | [25 empty] | 1
| 1 | 2 | < 10 | [25 empty] | 26
| 1 | 3 | < 10 | [25 empty] | 16
| 1 | 4 | < 10 | [25 empty] | < 10
| 1 | 5 | < 10 | [25 empty] | 32
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 550b (Certified Value 4.68 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 4.908 | 5.4 | 10.2584 | 5.7429 | 4.883 | 3.14 | 4.04 | [19 empty] | 1
| 1 | 2 | 4.873 | 5.6 | 10.2888 | 5.5215 | 4.812 | 3.07 | 4.11 | [19 empty] | 14
| 1 | 3 | 4.911 | 5.4 | 10.2928 | 5.6377 | 4.739 | 3.21 | 4.31 | [19 empty] | 16
| 1 | 4 | 4.987 | 5.6 | 10.2944 | 5.485 | 4.956 | 3.26 | 4.21 | [19 empty] | 4.6782
| 1 | 5 | 4.882 | 5.3 | 10.2504 | 5.2882 | 4.914 | 3.31 | 4.11 | [19 empty] | 185
| 1 | 6 | 4.827 | 5.5 | 10.2376 | 5.2644 | 4.814 | 3.21 | 4.19
| Mean |  | 4.898 | 5.4667 | 10.2704 | 5.4899 | 4.853 | 3.2 | 4.1617
| Median |  | 4.895 | 5.45 | 10.2736 | 5.5032 | 4.8485 | 3.21 | 4.15
| Std Dev. |  | 0.053126 | 0.12111 | 0.024642 | 0.18889 | 0.079282 | 0.085323 | 0.095167
| Rel.Std.Dev. |  | 1.085% | 2.215% | 0.2399% | 3.441% | 1.634% | 2.666% | 2.287%
| PDM3 |  | 4.698% | 16.85% | 119.5% | 17.35% | 3.736% | -31.6% | -11.04%
| Z-Score (Absolute) |  | 0 | 0.65 | 6.12 | 0.67 | 0.05 | 1.93 | 0.84
| Analytical results for Pt in OREAS 550b (Indicative Value 1.67 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 09 | 25 | [24 empty] | ppb
|  |  | AR*OES/MS | AR*OES/MS | [24 empty] | 2
|  |  | 01g | 0.1g | [24 empty] | 2
| 1 | 1 | < 5 | 3 | [24 empty] | 1
| 1 | 2 | < 5 | < 1 | [24 empty] | 2
| 1 | 3 | < 5 | < 1 | [24 empty] | 16
| 1 | 4 | < 5 | 2 | [24 empty] | 1.6667
| 1 | 5 | < 5 | 2 | [24 empty] | 30
| 1 | 6 | < 5 | 2
| Mean |  |  | 2.25
| Median |  |  | 2
| Std Dev. |  |  | 0.5
| Rel.Std.Dev. |  |  | 22.22%
| PDM3 |  |  | 35%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 550b (Certified Value 8.7 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 2
| 1 | 1 | 9.68 | 9.4 | 13.6 | 9.2 | 7.9 | 9.3893 | 7.8 | 7.7 | 8 | 9.8 | 8 | 9.8901 | 7.7 | 8.1 | 8.7 | 8.9 | 10.783 | 5.78 | 21 | 7.38 | [6 empty] | 1
| 1 | 2 | 9.49 | 10.8 | 13.7 | 9.5 | 8.4 | 8.7397 | 7.8 | 8.1 | 8.1 | 10.1 | 8.2 | 10.1102 | 8.1 | 8.6 | 8.9 | 9.4 | 10.495 | 5.79 | 18 | 6.97 | [6 empty] | 15
| 1 | 3 | 9.75 | 10.6 | 14.5 | 9.3 | 7.7 | 8.4751 | 7.6 | 8.1 | 8.4 | 9.9 | 8.2 | 10.357 | 7.9 | 8.6 | 8.8 | 9.1 | 10.429 | 5.96 | 20.3 | 7.72 | [6 empty] | 16
| 1 | 4 | 9.87 | 10.6 | 14.5 | 9.4 | 8.6 | 8.8099 | 7.8 | 7.9 | 8 | 9.6 | 8.2 | 10.3215 | 8.1 | 8.5 | 8.7 | 9.2 | 10.715 | 6.06 | 21.3 | 7.18 | [6 empty] | 8.7004
| 1 | 5 | 9.67 | 10.3 | 14.7 | 9.3 | 7.8 | 9.4599 | 7.7 | 7.7 | 8.2 | 9.7 | 8.4 | 10.4574 | 8.3 | 8.4 | 8.9 | 8.9 | 10.563 | 6.2 | 20.7 | 7.03 | [6 empty] | 186
| 1 | 6 | 9.67 | 10.3 | 14.4 | 9.6 | 8.5 | 8.5409 | 7.7 | 7.9 | 8.3 | 9.5 | 8.9 | 10.0536 | 8.1 | 8.6 | 8.7 | 8.9 | 10.346 | 5.96 | 21.2 | 6.95
| Mean |  | 9.6883 | 10.3333 | 14.2333 | 9.3833 | 8.15 | 8.9025 | 7.7333 | 7.9 | 8.1667 | 9.7667 | 8.3167 | 10.1983 | 8.0333 | 8.4667 | 8.7833 | 9.0667 | 10.5552 | 5.9583 | 20.4167 | 7.205
| Median |  | 9.675 | 10.45 | 14.45 | 9.35 | 8.15 | 8.7748 | 7.75 | 7.9 | 8.15 | 9.75 | 8.2 | 10.2158 | 8.1 | 8.55 | 8.75 | 9 | 10.529 | 5.96 | 20.85 | 7.105
| Std Dev. |  | 0.124 | 0.49666 | 0.46332 | 0.1472 | 0.3937 | 0.42336 | 0.08165 | 0.17889 | 0.1633 | 0.21602 | 0.31252 | 0.21502 | 0.20656 | 0.19664 | 0.098319 | 0.20656 | 0.16781 | 0.16055 | 1.2384 | 0.29912
| Rel.Std.Dev. |  | 1.28% | 4.806% | 3.255% | 1.569% | 4.831% | 4.756% | 1.056% | 2.264% | 2% | 2.212% | 3.758% | 2.108% | 2.571% | 2.323% | 1.119% | 2.278% | 1.59% | 2.695% | 6.066% | 4.151%
| PDM3 |  | 11.35% | 18.77% | 63.59% | 7.849% | -6.326% | 2.322% | -11.12% | -9.2% | -6.135% | 12.26% | -4.411% | 17.22% | -7.667% | -2.687% | 0.9529% | 4.209% | 21.32% | -31.52% | 134.7% | -17.19%
| Z-Score (Absolute) |  | 0.64 | 1.14 | 4.11 | 0.41 | 0.53 | 0.05 | 0.85 | 0.72 | 0.52 | 0.7 | 0.4 | 1.03 | 0.62 | 0.29 | 0.05 | 0.17 | 1.31 | 2.2 | 8.82 | 1.25
| Analytical results for Re in OREAS 550b (Certified Value 0.012 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 21 | 25 | 26 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.1g | 0.4g | [11 empty] | 3
| 1 | 1 | 0.01 | 0.011 | < 0.05 | 0.003 | 0.013 | < 0.05 | 0.013 | 0.012 | 0.01 | 0.011 | 0.012 | 0.016782 | 0.011 | 0.012 | < 0.05 | [11 empty] | 1
| 1 | 2 | 0.012 | 0.011 | < 0.05 | 0.003 | 0.011 | < 0.05 | 0.011 | 0.012 | 0.01 | 0.012 | 0.013 | 0.012299 | 0.012 | 0.012 | < 0.05 | [11 empty] | 16
| 1 | 3 | 0.012 | 0.01 | < 0.05 | 0.01 | 0.013 | < 0.05 | 0.013 | 0.012 | 0.012 | 0.013 | 0.012 | 0.013359 | 0.01 | 0.012 | < 0.05 | [11 empty] | 16
| 1 | 4 | 0.011 | 0.011 | < 0.05 | 0.001 | 0.011 | < 0.05 | 0.012 | 0.013 | 0.013 | 0.011 | 0.014 | 0.015755 | 0.012 | 0.012 | < 0.05 | [11 empty] | 0.011667
| 1 | 5 | 0.009 | 0.01 | < 0.05 | < 0.001 | 0.012 | < 0.05 | 0.012 | 0.011 | 0.011 | 0.013 | 0.012 | 0.013702 | 0.011 | 0.013 | < 0.05 | [11 empty] | 187
| 1 | 6 | 0.012 | 0.011 | < 0.05 | 0.005 | 0.011 | < 0.05 | 0.011 | 0.012 | 0.013 | 0.012 | 0.011 | 0.012191 | 0.01 | 0.013 | < 0.05
| Mean |  | 0.011 | 0.010667 |  | 0.0044 | 0.011833 |  | 0.012 | 0.012 | 0.0115 | 0.012 | 0.012333 | 0.014015 | 0.011 | 0.012333
| Median |  | 0.0115 | 0.011 |  | 0.003 | 0.0115 |  | 0.012 | 0.012 | 0.0115 | 0.012 | 0.012 | 0.013531 | 0.011 | 0.012
| Std Dev. |  | 0.0012649 | 0.0005164 |  | 0.0034351 | 0.00098319 |  | 0.00089443 | 0.00063246 | 0.0013784 | 0.00089443 | 0.0010328 | 0.0018702 | 0.00089443 | 0.0005164
| Rel.Std.Dev. |  | 11.5% | 4.841% |  | 78.07% | 8.309% |  | 7.454% | 5.27% | 11.99% | 7.454% | 8.374% | 13.34% | 8.131% | 4.187%
| PDM3 |  | -5.714% | -8.571% |  | -62.29% | 1.429% |  | 2.857% | 2.857% | -1.429% | 2.857% | 5.714% | 20.13% | -5.714% | 5.714%
| Z-Score (Absolute) |  | 0.67 | 0.9 | 8.77 | 5.56 | 0.11 | 8.77 | 0 | 0 | 0.34 | 0 | 0.22 | 1.36 | 0.67 | 0.22 | 8.77
| Analytical results for S in OREAS 550b (Certified Value 1.16 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 1.1471 | 1.08 | 1.107 | 1.19 | 1.1116 | 1.44 | 1.12 | 1.22 | 1.1567 | 1.16 | 1.22 | 1.18 | 1.28 | 1.17 | 1.2135 | 1.17 | 1.14 | 1.1 | 1.1491 | 1.11 | 1.13 | 1.3 | 1.05 | 1.16 | 1.14 | 1.165 | 1
| 1 | 2 | 1.1674 | 1.08 | 1.159 | 1.19 | 1.1164 | 1.37 | 1.11 | 1.21 | 1.1421 | 1.18 | 1.25 | 1.17 | 1.3 | 1.14 | 1.1822 | 1.2 | 1.17 | 1.05 | 1.1641 | 1.12 | 1.14 | 1.31 | 1.046 | 1.13 | 1.2 | 1.169 | 30
| 1 | 3 | 1.1119 | 1.11 | 1.172 | 1.19 | 1.1136 | 1.37 | 1.12 | 1.18 | 1.1551 | 1.18 | 1.24 | 1.16 | 1.28 | 1.15 | 1.1971 | 1.18 | 1.14 | 1.1 | 1.1457 | 1.08 | 1.14 | 1.28 | 1.057 | 1.18 | 1.22 | 1.179 | 16
| 1 | 4 | 1.1073 | 1.1 | 1.099 | 1.15 | 1.1098 | 1.37 | 1.12 | 1.22 | 1.1809 | 1.17 | 1.24 | 1.13 | 1.28 | 1.14 | 1.1914 | 1.2 | 1.16 | 1.11 | 1.1455 | 1.13 | 1.14 | 1.23 | 1.071 | 1.18 | 1.24 | 1.162 | 1.1612
| 1 | 5 | 1.1699 | 1.06 | 1.151 | 1.18 | 1.1097 | 1.39 | 1.11 | 1.17 | 1.1688 | 1.18 | 1.23 | 1.15 | 1.26 | 1.14 | 1.1772 | 1.22 | 1.17 | 1.12 | 1.1485 | 1.13 | 1.14 | 1.26 | 1.073 | 1.15 | 1.21 | 1.166 | 188
| 1 | 6 | 1.1463 | 1.1 | 1.09 | 1.19 | 1.1128 | 1.38 | 1.11 | 1.21 | 1.1283 | 1.16 | 1.26 | 1.14 | 1.26 | 1.14 | 1.1845 | 1.2 | 1.16 | 1.09 | 1.149 | 1.15 | 1.13 | 1.27 | 1.064 | 1.14 | 1.19 | 1.171
| Mean |  | 1.1417 | 1.0883 | 1.1297 | 1.1817 | 1.1123 | 1.3867 | 1.115 | 1.2017 | 1.1553 | 1.1717 | 1.24 | 1.155 | 1.2767 | 1.1467 | 1.191 | 1.195 | 1.1567 | 1.095 | 1.1503 | 1.12 | 1.1367 | 1.275 | 1.0602 | 1.1567 | 1.2 | 1.1687
| Median |  | 1.1467 | 1.09 | 1.129 | 1.19 | 1.1122 | 1.375 | 1.115 | 1.21 | 1.1559 | 1.175 | 1.24 | 1.155 | 1.28 | 1.14 | 1.1879 | 1.2 | 1.16 | 1.1 | 1.1488 | 1.125 | 1.14 | 1.275 | 1.0605 | 1.155 | 1.205 | 1.1675
| Std Dev. |  | 0.026749 | 0.018348 | 0.03503 | 0.016021 | 0.0025396 | 0.027325 | 0.0054772 | 0.02137 | 0.01866 | 0.0098319 | 0.014142 | 0.018708 | 0.015055 | 0.012111 | 0.013061 | 0.017607 | 0.013663 | 0.02429 | 0.0069427 | 0.023664 | 0.005164 | 0.02881 | 0.011053 | 0.020656 | 0.034059 | 0.0059554
| Rel.Std.Dev. |  | 2.343% | 1.686% | 3.101% | 1.356% | 0.2283% | 1.971% | 0.4912% | 1.778% | 1.615% | 0.8391% | 1.14% | 1.62% | 1.179% | 1.056% | 1.097% | 1.473% | 1.181% | 2.218% | 0.6036% | 2.113% | 0.4543% | 2.26% | 1.043% | 1.786% | 2.838% | 0.5096%
| PDM3 |  | -1.683% | -6.274% | -2.715% | 1.763% | -4.209% | 19.42% | -3.978% | 3.486% | -0.5063% | 0.9023% | 6.787% | -0.533% | 9.945% | -1.251% | 2.565% | 2.912% | -0.3895% | -5.7% | -0.9364% | -3.547% | -2.112% | 9.801% | -8.7% | -0.3895% | 3.342% | 0.6439%
| Z-Score (Absolute) |  | 0.27 | 1.29 | 0.5 | 0.49 | 0.83 | 4.38 | 0.78 | 0.87 | 0.01 | 0.3 | 1.6 | 0.02 | 2.29 | 0.18 | 0.66 | 0.74 | 0.01 | 1.16 | 0.11 | 0.68 | 0.37 | 2.26 | 1.82 | 0.01 | 0.84 | 0.24
| Analytical results for Sb in OREAS 550b (Certified Value 1.62 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 2
| 1 | 1 | 2.11 | 1.68 | 1.39 | 1.54 | 2.06 | 1.5 | 0.93954 | 1.58 | 1.76 | 1.73 | 1.47 | 1.62 | 1.628 | 1.46 | 1.69 | 1.28 | 1.31 | 1.4 | 1.69 | < 5 | 3.269 | 0.73 | 1.38 | 1.78 |  |  | 1
| 1 | 2 | 2.16 | 1.8 | 2.36 | 1.54 | 2 | 1.46 | 0.93954 | 1.6 | 1.76 | 1.58 | 1.54 | 1.54 | 1.5573 | 1.68 | 1.83 | 1.19 | 1.29 | 1.3 | 1.68 | < 5 | 2.22 | 0.72 | 1.55 | 1.76 |  |  | 17
| 1 | 3 | 2.2 | 1.77 | 2.06 | 1.53 | 1.94 | 1.44 | 0.97474 | 1.61 | 1.82 | 1.62 | 1.51 | 1.58 | 1.5829 | 1.5 | 1.79 | 1.19 | 1.29 | 1.4 | 1.6 | < 5 | 2.051 | 0.72 | 1.25 | 1.85 |  |  | 16
| 1 | 4 | 2.19 | 1.83 | 2.03 | 1.57 | 1.79 | 1.49 | 0.93951 | 1.62 | 1.78 | 1.54 | 1.43 | 1.64 | 1.6865 | 1.68 | 1.74 | 1.22 | 1.34 | 1.3 | 1.65 | < 5 | 2.048 | 0.8 | 1.59 | 1.81 |  |  | 1.6192
| 1 | 5 | 2.11 | 1.62 | 2.01 | 1.45 | 2.02 | 1.43 | 0.97626 | 1.57 | 1.76 | 1.65 | 1.5 | 1.64 | 1.5457 | 1.67 | 1.77 | 1.23 | 1.31 | 1.3 | 1.62 | < 5 | 2.319 | 0.78 | 1.59 | 1.8 |  |  | 189
| 1 | 6 | 2.13 | 1.7 | 1.57 | 1.55 | 1.76 | 1.48 | 0.9619 | 1.55 | 1.84 | 1.5 | 1.48 | 1.65 | 1.5647 | 1.64 | 1.84 | 1.27 | 1.4 | 1.3 | 1.64 | < 5 | 1.9 | 0.82 | 1.2 | 2
| Mean |  | 2.15 | 1.7333 | 1.9033 | 1.53 | 1.9283 | 1.4667 | 0.95525 | 1.5883 | 1.7867 | 1.6033 | 1.4883 | 1.6117 | 1.5942 | 1.605 | 1.7767 | 1.23 | 1.3233 | 1.3333 | 1.6467 |  | 2.3012 | 0.76167 | 1.4267 | 1.8333
| Median |  | 2.145 | 1.735 | 2.02 | 1.54 | 1.97 | 1.47 | 0.95072 | 1.59 | 1.77 | 1.6 | 1.49 | 1.63 | 1.5738 | 1.655 | 1.78 | 1.225 | 1.31 | 1.3 | 1.645 |  | 2.1355 | 0.755 | 1.465 | 1.805
| Std Dev. |  | 0.039497 | 0.079917 | 0.35641 | 0.041473 | 0.12529 | 0.028048 | 0.017927 | 0.026394 | 0.035024 | 0.082138 | 0.037639 | 0.043089 | 0.053611 | 0.098742 | 0.056451 | 0.038471 | 0.041793 | 0.05164 | 0.034448 |  | 0.49623 | 0.044008 | 0.17512 | 0.087101
| Rel.Std.Dev. |  | 1.837% | 4.611% | 18.73% | 2.711% | 6.497% | 1.912% | 1.877% | 1.662% | 1.96% | 5.123% | 2.529% | 2.674% | 3.363% | 6.152% | 3.177% | 3.128% | 3.158% | 3.873% | 2.092% |  | 21.56% | 5.778% | 12.27% | 4.751%
| PDM3 |  | 32.78% | 7.05% | 17.55% | -5.508% | 19.09% | -9.419% | -41% | -1.905% | 10.34% | -0.9789% | -8.081% | -0.4643% | -1.543% | -0.876% | 9.726% | -24.04% | -18.27% | -17.65% | 1.697% |  | 42.12% | -52.96% | -11.89% | 13.23%
| Z-Score (Absolute) |  | 2.04 | 0.48 | 1.12 | 0.28 | 1.21 | 0.52 | 2.43 | 0.06 | 0.68 | 0 | 0.43 | 0.03 | 0.04 | 0 | 0.65 | 1.4 | 1.05 | 1.01 | 0.16 | 3.36 | 2.61 | 3.16 | 0.66 | 0.86
| Analytical results for Sc in OREAS 550b (Certified Value 3.68 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  | 3
| 1 | 1 | 3.7 | 3.6 | 3.5 | 4 | 4.9 | 4 | 3.8 | 3.9131 | 3.8 | 3.8 | 3.9 | 3.9 | 3.5 | 3.9026 | 2.9 | 3.7 | 3.2 | 3.6 | 3.4 | 3.492 | 3.1 | 4 | 3.9 |  |  |  | 1
| 1 | 2 | 3.8 | 3.5 | 3.6 | 4 | 4.6 | 3.8 | 3.5 | 3.7926 | 3.8 | 3.9 | 3.9 | 3.9 | 3.5 | 3.7306 | 3.2 | 3.9 | 3.3 | 3.8 | 3.4 | 3.478 | 3 | 4 | 4 |  |  |  | 18
| 1 | 3 | 3.6 | 3.6 | 3.7 | 4 | 4.4 | 3.9 | 3.7 | 3.8868 | 3.7 | 3.9 | 4.1 | 3.9 | 3.5 | 3.9417 | 3.2 | 3.8 | 3.3 | 3.7 | 3.3 | 3.382 | 3.1 | 5 | 3.9 |  |  |  | 16
| 1 | 4 | 3.6 | 3.7 | 3.7 | 4 | 4.5 | 3.7 | 3.4 | 4.0124 | 3.8 | 3.7 | 3.9 | 3.7 | 3.6 | 3.8839 | 3.2 | 3.8 | 3.4 | 3.7 | 3.2 | 3.426 | 3.2 | 4 | 4 |  |  |  | 3.6846
| 1 | 5 | 3.7 | 3.6 | 3.6 | 4 | 4.8 | 4 | 3.9 | 3.9621 | 3.7 | 3.7 | 4.1 | 3.9 | 3.7 | 3.7786 | 3.4 | 3.9 | 3.4 | 3.7 | 3.3 | 3.491 | 3.2 | 4 | 4 |  |  |  | 190
| 1 | 6 | 3.7 | 3.3 | 3.7 | 4 | 4.2 | 3.7 | 3.6 | 3.7583 | 3.7 | 3.8 | 4.1 | 3.9 | 3.6 | 3.7709 | 3.4 | 3.9 | 3.5 | 3.7 | 3.4 | 3.468 | 3.2 | 5 | 4
| Mean |  | 3.6833 | 3.55 | 3.6333 | 4 | 4.5667 | 3.85 | 3.65 | 3.8876 | 3.75 | 3.8 | 4 | 3.8667 | 3.5667 | 3.8347 | 3.2167 | 3.8333 | 3.35 | 3.7 | 3.3333 | 3.4562 | 3.1333 | 4.3333 | 3.9667
| Median |  | 3.7 | 3.6 | 3.65 | 4 | 4.55 | 3.85 | 3.65 | 3.8999 | 3.75 | 3.8 | 4 | 3.9 | 3.55 | 3.8312 | 3.2 | 3.85 | 3.35 | 3.7 | 3.35 | 3.473 | 3.15 | 4 | 4
| Std Dev. |  | 0.075277 | 0.13784 | 0.08165 | 0 | 0.2582 | 0.13784 | 0.18708 | 0.097473 | 0.054772 | 0.089443 | 0.10954 | 0.08165 | 0.08165 | 0.085461 | 0.18348 | 0.08165 | 0.10488 | 0.063246 | 0.08165 | 0.043646 | 0.08165 | 0.5164 | 0.05164
| Rel.Std.Dev. |  | 2.044% | 3.883% | 2.247% | 0% | 5.654% | 3.58% | 5.126% | 2.507% | 1.461% | 2.354% | 2.739% | 2.112% | 2.289% | 2.229% | 5.704% | 2.13% | 3.131% | 1.709% | 2.449% | 1.263% | 2.606% | 11.92% | 1.302%
| PDM3 |  | -0.03508% | -3.654% | -1.392% | 8.559% | 23.94% | 4.488% | -0.9397% | 5.508% | 1.774% | 3.131% | 8.559% | 4.941% | -3.201% | 4.073% | -12.7% | 4.036% | -9.082% | 0.4172% | -9.534% | -6.2% | -14.96% | 17.61% | 7.655%
| Z-Score (Absolute) |  | 0.25 | 0.74 | 0.43 | 0.92 | 3 | 0.37 | 0.37 | 0.51 | 0 | 0.18 | 0.92 | 0.43 | 0.67 | 0.31 | 1.96 | 0.31 | 1.47 | 0.18 | 1.53 | 1.08 | 2.27 | 2.15 | 0.8
| Analytical results for Se in OREAS 550b (Certified Value 3.5 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 2
| 1 | 1 | 3 | 3.2 | 3.5 | 3 | < 5 | 4 | 3.3114 | 3.5 | 3.5 | 3.6 | 3.4 | 3.6 | 3.8668 | 4 | 3.9 | 3.4 | 3.7 | < 5 | 2.919 | < 0.1 | 3.39 | 3 |  |  |  |  | 1
| 1 | 2 | 3 | 3.6 | 3.4 | 3 | < 5 | 3 | 3.2631 | 3.8 | 3.5 | 3.3 | 3.5 | 3.6 | 4.2441 | 4 | 4.03 | 3.4 | 3.8 | < 5 | 3.109 | < 0.1 | 3.64 | 3 |  |  |  |  | 22
| 1 | 3 | 3 | 3.5 | 3.5 | 3 | < 5 | 4 | 3.5413 | 3.9 | 3.8 | 4.1 | 3.7 | 3.4 | 4.4265 | 4 | 4.08 | 3.1 | 4.2 | < 5 | 2.841 | < 0.1 | 3.24 | 3 |  |  |  |  | 16
| 1 | 4 | 4 | 3.3 | 3.4 | 3 | < 5 | 3 | 3.5066 | 3.9 | 4.1 | 3.6 | 3.5 | 3.6 | 3.8546 | 4 | 4.02 | 3.6 | 4 | < 5 | 3.078 | < 0.1 | 3.44 | 3 |  |  |  |  | 3.5038
| 1 | 5 | 3 | 3.5 | 3.1 | 3 | < 5 | 3 | 3.1038 | 3.8 | 3.7 | 3.6 | 3.5 | 3.6 | 3.7521 | 4 | 4.16 | 3.5 | 4.4 | < 5 | 2.908 | < 0.1 | 3.18 | 3 |  |  |  |  | 191
| 1 | 6 | 3 | 2.7 | 3.5 | 3 | < 5 | 4 | 3.0864 | 3.3 | 3.8 | 3.6 | 3.4 | 3.5 | 3.945 | 4 | 4.83 | 3.4 | 4.8 | < 5 | 3.01 | < 0.1 | 3.13 | 3
| Mean |  | 3.1667 | 3.3 | 3.4 | 3 |  | 3.5 | 3.3021 | 3.7 | 3.7333 | 3.6333 | 3.5 | 3.55 | 4.0149 | 4 | 4.17 | 3.4 | 4.15 |  | 2.9775 |  | 3.3367 | 3
| Median |  | 3 | 3.4 | 3.45 | 3 |  | 3.5 | 3.2872 | 3.8 | 3.75 | 3.6 | 3.5 | 3.6 | 3.9059 | 4 | 4.055 | 3.4 | 4.1 |  | 2.9645 |  | 3.315 | 3
| Std Dev. |  | 0.40825 | 0.32863 | 0.15492 | 0 |  | 0.54772 | 0.19311 | 0.24495 | 0.22509 | 0.2582 | 0.10954 | 0.083666 | 0.26212 | 0 | 0.33431 | 0.16733 | 0.40866 |  | 0.1052 |  | 0.19086 | 0
| Rel.Std.Dev. |  | 12.89% | 9.959% | 4.556% | 0% |  | 15.65% | 5.848% | 6.62% | 6.029% | 7.106% | 3.13% | 2.357% | 6.529% | 0% | 8.017% | 4.922% | 9.847% |  | 3.533% |  | 5.72% | 0%
| PDM3 |  | -9.622% | -5.817% | -2.963% | -14.38% |  | -0.1088% | -5.757% | 5.599% | 6.551% | 3.697% | -0.1088% | 1.318% | 14.59% | 14.16% | 19.01% | -2.963% | 18.44% |  | -15.02% |  | -4.77% | -14.38%
| Z-Score (Absolute) |  | 0.5 | 0.21 | 0 | 0.85 | 1.92 | 0.21 | 0.21 | 0.64 | 0.71 | 0.5 | 0.21 | 0.32 | 1.31 | 1.28 | 1.64 | 0 | 1.6 | 1.92 | 0.9 | 7.13 | 0.13 | 0.85
| Analytical results for Sm in OREAS 550b (Certified Value 3 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 3.166 | 3.7 | 6.141 | 3.5801 | 3.058 | 2.12 | 2.77 | [19 empty] | 1
| 1 | 2 | 3.13 | 3.5 | 6.1822 | 3.575 | 3.033 | 2.08 | 2.77 | [19 empty] | 5
| 1 | 3 | 3.143 | 2.7 | 6.1725 | 3.6076 | 3.074 | 2.16 | 2.86 | [19 empty] | 16
| 1 | 4 | 3.154 | 3.3 | 6.2006 | 3.5195 | 3.187 | 2.21 | 2.84 | [19 empty] | 3.0035
| 1 | 5 | 3.128 | 3.1 | 6.1734 | 3.3557 | 3.112 | 2.25 | 2.73 | [19 empty] | 192
| 1 | 6 | 3.11 | 3.6 | 6.1833 | 3.4205 | 3.083 | 2.16 | 2.84
| Mean |  | 3.1385 | 3.3167 | 6.1755 | 3.5097 | 3.0912 | 2.1633 | 2.8017
| Median |  | 3.1365 | 3.4 | 6.1778 | 3.5473 | 3.0785 | 2.16 | 2.805
| Std Dev. |  | 0.020057 | 0.37103 | 0.01972 | 0.10058 | 0.053783 | 0.060882 | 0.051929
| Rel.Std.Dev. |  | 0.6391% | 11.19% | 0.3193% | 2.866% | 1.74% | 2.814% | 1.854%
| PDM3 |  | 4.494% | 10.43% | 105.6% | 16.85% | 2.918% | -27.97% | -6.72%
| Z-Score (Absolute) |  | 0 | 0.36 | 6.08 | 0.74 | 0.09 | 1.95 | 0.67
| Analytical results for Sn in OREAS 550b (Certified Value 0.83 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 3
| 1 | 1 | 0.83 | 0.84 | 1 | 0.8 | 0.9 | 0.80858 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.88511 | 0.79 | 0.8 | 0.8 | < 5 | 0.959 | 0.56 | 0.9 | 0.9 | [6 empty] | 1
| 1 | 2 | 0.78 | 0.86 | 1.1 | 0.9 | 0.8 | 0.80739 | 0.8 | 0.8 | 0.8 | 0.9 | 0.8 | 0.85725 | 0.82 | 0.8 | 0.8 | < 5 | 0.933 | 0.56 | 0.9 | 0.9 | [6 empty] | 6
| 1 | 3 | 0.81 | 0.88 | 1 | 0.8 | 0.9 | 0.79619 | 0.8 | 0.8 | 0.8 | 0.9 | 0.8 | 0.86221 | 0.84 | 0.8 | 0.8 | < 5 | 0.917 | 0.57 | 0.9 | 0.9 | [6 empty] | 16
| 1 | 4 | 0.8 | 0.83 | 1 | 0.8 | 0.8 | 0.79704 | 0.8 | 0.8 | 1 | 0.8 | 0.8 | 0.85772 | 0.83 | 0.8 | 0.8 | < 5 | 0.938 | 0.59 | 0.8 | 0.9 | [6 empty] | 0.83287
| 1 | 5 | 0.84 | 0.86 | 1.1 | 0.8 | 0.9 | 0.79309 | 0.8 | 0.8 | 0.8 | 0.9 | 0.8 | 0.8497 | 0.87 | 0.8 | 0.8 | < 5 | 0.809 | 0.58 | 0.8 | 0.9 | [6 empty] | 193
| 1 | 6 | 0.81 | 0.82 | 1.2 | 0.9 | 0.8 | 0.80647 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.87134 | 0.87 | 0.8 | 0.8 | < 5 | 0.825 | 0.57 | 0.9 | 0.9
| Mean |  | 0.81167 | 0.84833 | 1.0667 | 0.83333 | 0.85 | 0.80146 | 0.8 | 0.8 | 0.83333 | 0.85 | 0.8 | 0.86389 | 0.83667 | 0.8 | 0.8 |  | 0.89683 | 0.57167 | 0.86667 | 0.9
| Median |  | 0.81 | 0.85 | 1.05 | 0.8 | 0.85 | 0.80176 | 0.8 | 0.8 | 0.8 | 0.85 | 0.8 | 0.85996 | 0.835 | 0.8 | 0.8 |  | 0.925 | 0.57 | 0.9 | 0.9
| Std Dev. |  | 0.02137 | 0.022286 | 0.08165 | 0.05164 | 0.054772 | 0.0067568 | 1.2162e-16 | 1.2162e-16 | 0.08165 | 0.054772 | 1.2162e-16 | 0.012586 | 0.030768 | 1.2162e-16 | 1.2162e-16 |  | 0.06348 | 0.01169 | 0.05164 | 0
| Rel.Std.Dev. |  | 2.633% | 2.627% | 7.655% | 6.197% | 6.444% | 0.8431% | 1.52e-14% | 1.52e-14% | 9.798% | 6.444% | 1.52e-14% | 1.457% | 3.677% | 1.52e-14% | 1.52e-14% |  | 7.078% | 2.045% | 5.958% | 0%
| PDM3 |  | -2.546% | 1.856% | 28.07% | 0.05519% | 2.056% | -3.771% | -3.947% | -3.947% | 0.05519% | 2.056% | -3.947% | 3.724% | 0.4554% | -3.947% | -3.947% |  | 7.679% | -31.36% | 4.057% | 8.06%
| Z-Score (Absolute) |  | 0.46 | 0.26 | 4.56 | 0.03 | 0.3 | 0.66 | 0.69 | 0.69 | 0.03 | 0.3 | 0.69 | 0.57 | 0.03 | 0.69 | 0.69 | 32.76 | 1.22 | 5.18 | 0.62 | 1.28
| Analytical results for Sr in OREAS 550b (Certified Value 29.6 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 2
| 1 | 1 | 31.36 | 29.4 | 34.7 | 33.5 | 37.3 | 30 | 26.7 | 31.6089 | 31 | 27.9 | 28 | 32.3 | 28.5 | 31.9153 | 25.9 | 29.7 | 40.9 | 27 | 29 | 24.5 | 28.041 | 39 | 32.4 | 27.3 |  |  | 1
| 1 | 2 | 31.15 | 30.2 | 31.7 | 33.2 | 39.1 | 30 | 29 | 30.6174 | 30.8 | 29.3 | 28.8 | 32.5 | 28.7 | 31.1979 | 25.8 | 29.8 | 39.8 | 27 | 28 | 24.8 | 26.849 | 38 | 27.2 | 27.5 |  |  | #N/A
| 1 | 3 | 31.56 | 31.7 | 31.6 | 33.3 | 33.4 | 30 | 28.2 | 31.1693 | 30.1 | 29.5 | 29.6 | 32.4 | 27.8 | 31.1305 | 26 | 30.2 | 40.2 | 26 | 28 | 24.5 | 26.691 | 38 | 29.6 | 27.9 |  |  | 16
| 1 | 4 | 31.6 | 32.2 | 34.1 | 33.7 | 32.4 | 31 | 27.2 | 32.1724 | 30.1 | 29.3 | 29.1 | 32 | 30 | 30.4975 | 26.4 | 30 | 40.5 | 28 | 28 | 24.6 | 27.858 | 39 | 30.1 | 27.4 |  |  | 29.6478
| 1 | 5 | 30.98 | 28.4 | 29.3 | 33.9 | 36.8 | 30 | 28.2 | 32.235 | 30.5 | 28.6 | 30.2 | 31.8 | 28.9 | 30.8282 | 25.9 | 29.6 | 38.2 | 27 | 28 | 22.1 | 27.335 | 37 | 26.3 | 27.1 |  |  | 194
| 1 | 6 | 30.7 | 28.6 | 34 | 33.9 | 31.1 | 30 | 29.3 | 31.2207 | 30.4 | 29.4 | 29 | 31.8 | 29.3 | 30.4045 | 25.8 | 29.9 | 39.5 | 29 | 28 | 25.6 | 26.892 | 40 | 30.6 | 27.5
| Mean |  | 31.225 | 30.0833 | 32.5667 | 33.5833 | 35.0167 | 30.1667 | 28.1 | 31.5039 | 30.4833 | 29 | 29.1167 | 32.1333 | 28.8667 | 30.9957 | 25.9667 | 29.8667 | 39.85 | 27.3333 | 28.1667 | 24.35 | 27.2777 | 38.5 | 29.3667 | 27.45
| Median |  | 31.255 | 29.8 | 32.85 | 33.6 | 35.1 | 30 | 28.2 | 31.4148 | 30.45 | 29.3 | 29.05 | 32.15 | 28.8 | 30.9794 | 25.9 | 29.85 | 40 | 27 | 28 | 24.55 | 27.1135 | 38.5 | 29.85 | 27.45
| Std Dev. |  | 0.34996 | 1.588 | 2.0646 | 0.29944 | 3.1581 | 0.40825 | 1.004 | 0.62777 | 0.3656 | 0.6261 | 0.74409 | 0.30768 | 0.74476 | 0.55325 | 0.22509 | 0.21602 | 0.94816 | 1.0328 | 0.40825 | 1.1777 | 0.56558 | 1.0488 | 2.2545 | 0.26646
| Rel.Std.Dev. |  | 1.121% | 5.279% | 6.34% | 0.8916% | 9.019% | 1.353% | 3.573% | 1.993% | 1.199% | 2.159% | 2.556% | 0.9575% | 2.58% | 1.785% | 0.8669% | 0.7233% | 2.379% | 3.779% | 1.449% | 4.837% | 2.073% | 2.724% | 7.677% | 0.9707%
| PDM3 |  | 5.32% | 1.469% | 9.845% | 13.27% | 18.11% | 1.75% | -5.221% | 6.261% | 2.818% | -2.185% | -1.792% | 8.383% | -2.635% | 4.546% | -12.42% | 0.7381% | 34.41% | -7.807% | -4.996% | -17.87% | -7.994% | 29.86% | -0.9483% | -7.413%
| Z-Score (Absolute) |  | 0.46 | 0.04 | 0.95 | 1.32 | 1.85 | 0.07 | 0.69 | 0.56 | 0.19 | 0.36 | 0.31 | 0.79 | 0.41 | 0.37 | 1.47 | 0.04 | 3.62 | 0.97 | 0.66 | 2.06 | 0.99 | 3.12 | 0.22 | 0.92
| Analytical results for Ta in OREAS 550b (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 24 | 25 | 26 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [8 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 7
| 1 | 3 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 195
| 1 | 6 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 550b (Certified Value 0.26 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 05 | 08 | 15 | 16 | 17 | 22 | 23 | 24 | 26 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | [15 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.4g | [15 empty] | 3
| 1 | 1 | 0.248 | 0.3 | 0.54808 | 0.26 | 0.29535 | 0.25 | 0.24 | < 5 | 0.247 | 0.19 | 0.27 | [15 empty] | 1
| 1 | 2 | 0.255 | 0.3 | 0.54384 | 0.23 | 0.28925 | 0.25 | 0.25 | < 5 | 0.252 | 0.19 | 0.27 | [15 empty] | 8
| 1 | 3 | 0.254 | 0.2 | 0.5816 | 0.26 | 0.29604 | 0.26 | 0.26 | < 5 | 0.252 | 0.19 | 0.28 | [15 empty] | 16
| 1 | 4 | 0.258 | 0.3 | 0.55856 | 0.24 | 0.29864 | 0.25 | 0.25 | < 5 | 0.252 | 0.2 | 0.26 | [15 empty] | 0.25977
| 1 | 5 | 0.259 | 0.3 | 0.54968 | 0.25 | 0.28342 | 0.25 | 0.25 | < 5 | 0.247 | 0.2 | 0.26 | [15 empty] | 196
| 1 | 6 | 0.243 | 0.2 | 0.58176 | 0.23 | 0.27736 | 0.27 | 0.25 | < 5 | 0.252 | 0.19 | 0.27
| Mean |  | 0.25283 | 0.26667 | 0.56059 | 0.245 | 0.29001 | 0.255 | 0.25 |  | 0.25033 | 0.19333 | 0.26833
| Median |  | 0.2545 | 0.3 | 0.55412 | 0.245 | 0.2923 | 0.25 | 0.25 |  | 0.252 | 0.19 | 0.27
| Std Dev. |  | 0.0061779 | 0.05164 | 0.017028 | 0.013784 | 0.0082866 | 0.0083666 | 0.0063246 |  | 0.002582 | 0.005164 | 0.0075277
| Rel.Std.Dev. |  | 2.443% | 19.36% | 3.038% | 5.626% | 2.857% | 3.281% | 2.53% |  | 1.031% | 2.671% | 2.805%
| PDM3 |  | -2.671% | 2.654% | 115.8% | -5.687% | 11.64% | -1.837% | -3.762% |  | -3.633% | -25.58% | 3.296%
| Z-Score (Absolute) |  | 0.13 | 0.67 | 17.66 | 0.58 | 2.02 | 0 | 0.29 | 129.76 | 0.27 | 3.56 | 0.77
| Analytical results for Te in OREAS 550b (Certified Value 0.048 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 3
| 1 | 1 | < 0.02 | 0.04 | < 0.05 | 0.15 | 0.07 | 0.06855 | 0.04 | 0.06 | 0.04 | 0.05 | 0.05 | < 0.2 | < 0.05 | 0.13 | < 0.2 | < 0.1 | < 5 | 0.085 | 0.03 | 0.05 | < 0.05 | [5 empty] | 1
| 1 | 2 | < 0.02 | 0.05 | < 0.05 | 0.15 | < 0.05 | 0.07032 | 0.02 | 0.04 | 0.04 | 0.05 | 0.06 | < 0.2 | < 0.05 | 0.08 | < 0.2 | < 0.1 | < 5 | 0.085 | 0.03 | 0.04 | 0.05 | [5 empty] | 9
| 1 | 3 | < 0.02 | 0.06 | < 0.05 | 0.16 | 0.05 | 0.062675 | 0.04 | 0.04 | 0.05 | 0.06 | 0.04 | < 0.2 | < 0.05 | 0.09 | < 0.2 | < 0.1 | < 5 | 0.099 | 0.03 | 0.04 | < 0.05 | [5 empty] | 16
| 1 | 4 | 0.03 | 0.06 | < 0.05 | 0.17 | 0.06 | 0.065165 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | < 0.2 | < 0.05 | 0.11 | < 0.2 | < 0.1 | < 5 | 0.1 | 0.04 | 0.04 | < 0.05 | [5 empty] | 0.048338
| 1 | 5 | < 0.02 | 0.06 | < 0.05 | 0.17 | 0.07 | 0.057848 | 0.03 | 0.06 | 0.03 | 0.04 | 0.04 | < 0.2 | < 0.05 | 0.1 | < 0.2 | < 0.1 | < 5 | 0.033 | 0.03 | 0.04 | < 0.05 | [5 empty] | 197
| 1 | 6 | 0.04 | 0.03 | 0.05 | 0.15 | < 0.05 | 0.057617 | 0.03 | 0.06 | 0.01 | 0.05 | 0.05 | < 0.2 | < 0.05 | 0.45 | < 0.2 | < 0.1 | < 5 | 0.059 | 0.04 | 0.05 | < 0.05
| Mean |  | 0.035 | 0.05 | 0.05 | 0.15833 | 0.0625 | 0.063696 | 0.033333 | 0.051667 | 0.033333 | 0.05 | 0.048333 |  |  | 0.16 |  |  |  | 0.076833 | 0.033333 | 0.043333 | 0.05
| Median |  | 0.035 | 0.055 | 0.05 | 0.155 | 0.065 | 0.06392 | 0.035 | 0.055 | 0.035 | 0.05 | 0.05 |  |  | 0.105 |  |  |  | 0.085 | 0.03 | 0.04 | 0.05
| Std Dev. |  | 0.0070711 | 0.012649 |  | 0.0098319 | 0.0095743 | 0.0053255 | 0.008165 | 0.0098319 | 0.013663 | 0.0063246 | 0.0075277 |  |  | 0.14311 |  |  |  | 0.02608 | 0.005164 | 0.005164
| Rel.Std.Dev. |  | 20.2% | 25.3% |  | 6.21% | 15.32% | 8.361% | 24.49% | 19.03% | 40.99% | 12.65% | 15.57% |  |  | 89.44% |  |  |  | 33.94% | 15.49% | 11.92%
| PDM3 |  | -27.59% | 3.439% | 3.439% | 227.6% | 29.3% | 31.77% | -31.04% | 6.887% | -31.04% | 3.439% | -0.009238% |  |  | 231% |  |  |  | 58.95% | -31.04% | -10.35% | 3.439%
| Z-Score (Absolute) |  | 1.28 | 0 | 0.84 | 4.38 | 0 | 0.55 | 0.67 | 0.07 | 0.67 | 0 | 0.07 | 2.02 | 1.01 | 4.45 | 2.02 | 0 | 99.12 | 1.09 | 0.67 | 0.27 | 0.84
| Analytical results for Th in OREAS 550b (Certified Value 6.1 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 3
| 1 | 1 | 6.39 | 5.8 | 7 | 6.6 | 5.8 | 6.1 | 6.5746 | 5.8 | 5.3 | 5.1 | 7.2 | 5.6 | 6.272 | 5.7 | 5.4 | 6.2 | 6.8 | 5.475 | 3.64 | 7.17 | 5.7 | [5 empty] | 1
| 1 | 2 | 6.28 | 6 | 6.5 | 7.08 | 5.8 | 6.2 | 6.4534 | 5.7 | 5.6 | 5.1 | 7.2 | 5.5 | 6.2843 | 6 | 5.5 | 6.2 | 7.1 | 5.423 | 3.52 | 6.6 | 5.66 | [5 empty] | 10
| 1 | 3 | 6.46 | 6.3 | 6.2 | 7.18 | 5.8 | 5.9 | 6.494 | 5.7 | 5.8 | 5.3 | 7 | 5.4 | 6.5551 | 6.1 | 5.3 | 6.1 | 7.2 | 5.36 | 3.64 | 7.01 | 5.63 | [5 empty] | 16
| 1 | 4 | 6.46 | 6.3 | 6.9 | 7.12 | 5.7 | 6.2 | 6.698 | 5.8 | 5.7 | 5.2 | 6.6 | 5.9 | 6.4182 | 5.8 | 5.5 | 6 | 6.3 | 5.596 | 3.69 | 7.34 | 5.74 | [5 empty] | 6.1023
| 1 | 5 | 6.35 | 5.6 | 6 | 7.07 | 6 | 5.9 | 6.7497 | 5.7 | 5.5 | 5.3 | 6.9 | 5.8 | 6.381 | 5.6 | 5.7 | 6.1 | 6.5 | 5.577 | 3.76 | 6.11 | 5.54 | [5 empty] | 198
| 1 | 6 | 6.31 | 5.2 | 6.6 | 7.16 | 5.6 | 6.1 | 6.5499 | 5.7 | 5.7 | 5.2 | 6.9 | 5.7 | 6.3853 | 6 | 5.8 | 6 | 6.6 | 5.418 | 3.66 | 6.94 | 5.81
| Mean |  | 6.375 | 5.8667 | 6.5333 | 7.035 | 5.7833 | 6.0667 | 6.5866 | 5.7333 | 5.6 | 5.2 | 6.9667 | 5.65 | 6.3826 | 5.8667 | 5.5333 | 6.1 | 6.75 | 5.4748 | 3.6517 | 6.8617 | 5.68
| Median |  | 6.37 | 5.9 | 6.55 | 7.1 | 5.8 | 6.1 | 6.5622 | 5.7 | 5.65 | 5.2 | 6.95 | 5.65 | 6.3831 | 5.9 | 5.5 | 6.1 | 6.7 | 5.449 | 3.65 | 6.975 | 5.68
| Std Dev. |  | 0.075565 | 0.4274 | 0.38816 | 0.21742 | 0.13292 | 0.13663 | 0.1156 | 0.05164 | 0.17889 | 0.089443 | 0.22509 | 0.18708 | 0.1029 | 0.19664 | 0.18619 | 0.089443 | 0.35071 | 0.094046 | 0.078592 | 0.44396 | 0.093167
| Rel.Std.Dev. |  | 1.185% | 7.285% | 5.941% | 3.09% | 2.298% | 2.252% | 1.755% | 0.9007% | 3.194% | 1.72% | 3.231% | 3.311% | 1.612% | 3.352% | 3.365% | 1.466% | 5.196% | 1.718% | 2.152% | 6.47% | 1.64%
| PDM3 |  | 4.469% | -3.861% | 7.064% | 15.28% | -5.227% | -0.5837% | 7.937% | -6.046% | -8.231% | -14.79% | 14.16% | -7.412% | 4.594% | -3.861% | -9.324% | -0.03747% | 10.61% | -10.28% | -40.16% | 12.44% | -6.92%
| Z-Score (Absolute) |  | 0.87 | 0 | 1.15 | 2.01 | 0.14 | 0.34 | 1.24 | 0.23 | 0.46 | 1.15 | 1.89 | 0.37 | 0.89 | 0 | 0.57 | 0.4 | 1.52 | 0.67 | 3.81 | 1.71 | 0.32
| Analytical results for Ti in OREAS 550b (Indicative Value 0.003 wt.%) | [27 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  | 3
| 1 | 1 | 0.002 | 0.001 | < 0.01 | 0.0031 | < 0.005 | < 0.01 | 0.0031139 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.0026975 | < 0.01 | < 0.01 | 0.0047 | 0.002 | < 0.005 | < 0.01 | 0.0043714 | 0.0036 | 0.0076 | 0.003 |  |  |  | 1
| 1 | 2 | 0.002 | 0.002 | < 0.01 | 0.003 | < 0.005 | < 0.01 | 0.0030471 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.0026571 | < 0.01 | < 0.01 | 0.0047 | 0.002 | < 0.005 | < 0.01 | 0.0044234 | 0.0048 | 0.0063 | 0.004 |  |  |  | 25
| 1 | 3 | 0.0019 | 0.002 | < 0.01 | 0.0032 | < 0.005 | < 0.01 | 0.0031574 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.0026166 | < 0.01 | < 0.01 | 0.0046 | 0.002 | < 0.005 | < 0.01 | 0.0043554 | 0.003 | 0.0065 | 0.003 |  |  |  | 16
| 1 | 4 | 0.0019 | 0.002 | < 0.01 | 0.003 | < 0.005 | < 0.01 | 0.003115 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.0026621 | < 0.01 | < 0.01 | 0.0044 | 0.002 | < 0.005 | < 0.01 | 0.004578 | 0.0018 | 0.007 | 0.003 |  |  |  | 0.0034021
| 1 | 5 | 0.002 | 0.001 | < 0.01 | 0.0031 | < 0.005 | < 0.01 | 0.0031598 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.0025911 | < 0.01 | < 0.01 | 0.0047 | 0.002 | < 0.005 | < 0.01 | 0.0045401 | 0.0042 | 0.007 | 0.003 |  |  |  | 31
| 1 | 6 | 0.002 | 0.002 | < 0.01 | 0.0031 | < 0.005 | < 0.01 | 0.0028611 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.0026035 | < 0.01 | < 0.01 | 0.0048 | 0.002 | < 0.005 | < 0.01 | 0.0043475 | 0.0042 | 0.0079 | 0.003
| Mean |  | 0.0019667 | 0.0016667 |  | 0.0030833 |  |  | 0.0030757 | [5 empty] | 0.002638 |  |  | 0.00465 | 0.002 |  |  | 0.004436 | 0.0036 | 0.00705 | 0.0031667
| Median |  | 0.002 | 0.002 |  | 0.0031 |  |  | 0.0031145 | [5 empty] | 0.0026368 |  |  | 0.0047 | 0.002 |  |  | 0.0043974 | 0.0039 | 0.007 | 0.003
| Std Dev. |  | 5.164e-05 | 0.0005164 |  | 7.5277e-05 |  |  | 0.00011281 | [5 empty] | 4.0803e-05 |  |  | 0.00013784 | 0 |  |  | 9.9661e-05 | 0.0010733 | 0.00061563 | 0.00040825
| Rel.Std.Dev. |  | 2.626% | 30.98% |  | 2.441% |  |  | 3.668% | [5 empty] | 1.547% |  |  | 2.964% | 0% |  |  | 2.247% | 29.81% | 8.732% | 12.89%
| PDM3 |  | -42.19% | -51.01% |  | -9.369% |  |  | -9.593% | [5 empty] | -22.46% |  |  | 36.68% | -41.21% |  |  | 30.39% | 5.818% | 107.2% | -6.92%
| Z-Score (Absolute) |  | 1.3 | 1.65 | 2.25 | 0.01 | 0.67 | 2.25 | 0 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.51 | 2.25 | 2.25 | 1.84 | 1.26 | 0.67 | 2.25 | 1.59 | 0.61 | 4.66 | 0.11
| Analytical results for Tl in OREAS 550b (Indicative Value 0.18 ppm) | [27 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.1g | 0.4g |  |  |  | 2
| 1 | 1 | 0.24 | 0.17 | < 0.02 | 0.26 | 0.1 | 0.17 | < 0.05 | 0.1 | 0.12 | 0.13 | 0.12 | 0.11 | 0.14301 | 0.2 | 0.13 | 0.29 | 0.247 | 0.2 | < 0.05 | < 5 | 0.691 | 0.31 | 0.15 |  |  |  | 1
| 1 | 2 | 0.24 | 0.17 | 0.02 | 0.26 | 0.1 | 0.19 | < 0.05 | 0.11 | 0.13 | 0.15 | 0.12 | 0.12 | 0.15051 | 0.2 | 0.14 | 0.3 | 0.239 | 0.2 | < 0.05 | < 5 | 0.396 | 0.3 | 0.15 |  |  |  | 26
| 1 | 3 | 0.24 | 0.17 | < 0.02 | 0.27 | 0.11 | 0.18 | < 0.05 | 0.12 | 0.12 | 0.12 | 0.12 | 0.11 | 0.14683 | 0.2 | 0.12 | 0.29 | 0.246 | 0.2 | < 0.05 | < 5 | 0.333 | 0.3 | 0.16 |  |  |  | 16
| 1 | 4 | 0.24 | 0.18 | < 0.02 | 0.26 | 0.1 | 0.19 | < 0.05 | 0.12 | 0.11 | 0.12 | 0.11 | 0.12 | 0.14502 | 0.18 | 0.12 | 0.31 | 0.263 | 0.2 | < 0.05 | < 5 | 0.309 | 0.31 | 0.15 |  |  |  | 0.17627
| 1 | 5 | 0.23 | 0.17 | < 0.02 | 0.26 | 0.1 | 0.17 | < 0.05 | 0.11 | 0.11 | 0.13 | 0.11 | 0.12 | 0.13638 | 0.19 | 0.12 | 0.28 | 0.253 | 0.2 | < 0.05 | < 5 | 0.28 | 0.27 | 0.15 |  |  |  | 32
| 1 | 6 | 0.23 | 0.2 | < 0.02 | 0.27 | 0.1 | 0.18 | < 0.05 | 0.11 | 0.11 | 0.13 | 0.13 | 0.12 | 0.15479 | 0.18 | 0.12 | 0.28 | 0.266 | 0.2 | < 0.05 | < 5 | 0.271 | 0.32 | 0.15
| Mean |  | 0.23667 | 0.17667 | 0.02 | 0.26333 | 0.10167 | 0.18 |  | 0.11167 | 0.11667 | 0.13 | 0.11833 | 0.11667 | 0.14609 | 0.19167 | 0.125 | 0.29167 | 0.25233 | 0.2 |  |  | 0.38 | 0.30167 | 0.15167
| Median |  | 0.24 | 0.17 | 0.02 | 0.26 | 0.1 | 0.18 |  | 0.11 | 0.115 | 0.13 | 0.12 | 0.12 | 0.14592 | 0.195 | 0.12 | 0.29 | 0.25 | 0.2 |  |  | 0.321 | 0.305 | 0.15
| Std Dev. |  | 0.005164 | 0.012111 |  | 0.005164 | 0.0040825 | 0.0089443 |  | 0.0075277 | 0.008165 | 0.010954 | 0.0075277 | 0.005164 | 0.0063345 | 0.0098319 | 0.0083666 | 0.01169 | 0.010463 | 3.0405e-17 |  |  | 0.15881 | 0.017224 | 0.0040825
| Rel.Std.Dev. |  | 2.182% | 6.855% |  | 1.961% | 4.016% | 4.969% |  | 6.741% | 6.999% | 8.427% | 6.361% | 4.426% | 4.336% | 5.13% | 6.693% | 4.008% | 4.146% | 1.52e-14% |  |  | 41.79% | 5.71% | 2.692%
| PDM3 |  | 34.27% | 0.2267% | -88.65% | 49.39% | -42.32% | 2.118% |  | -36.65% | -33.81% | -26.25% | -32.87% | -33.81% | -17.12% | 8.737% | -29.08% | 65.47% | 43.15% | 13.46% |  |  | 115.6% | 71.14% | -13.96%
| Z-Score (Absolute) |  | 1.19 | 0.35 | 1.95 | 1.56 | 0.7 | 0.4 | 1.77 | 0.56 | 0.49 | 0.3 | 0.47 | 0.49 | 0.08 | 0.56 | 0.37 | 1.95 | 1.4 | 0.67 | 1.77 | 32.76 | 3.19 | 2.09 | 0
| Analytical results for Tm in OREAS 550b (Indicative Value 0.071 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [21 empty] | 3
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | [21 empty] | 3
| 1 | 1 | 0.067 | < 0.1 | 0.13224 | 0.07691 | 0.066 | [21 empty] | 1
| 1 | 2 | 0.074 | < 0.1 | 0.13248 | 0.075019 | 0.062 | [21 empty] | 12
| 1 | 3 | 0.07 | < 0.1 | 0.1376 | 0.077009 | 0.065 | [21 empty] | 16
| 1 | 4 | 0.071 | < 0.1 | 0.14792 | 0.079505 | 0.068 | [21 empty] | 0.071142
| 1 | 5 | 0.073 | < 0.1 | 0.13872 | 0.076193 | 0.066 | [21 empty] | 30
| 1 | 6 | 0.068 | < 0.1 | 0.13496 | 0.069656 | 0.069
| Mean |  | 0.0705 |  | 0.13732 | 0.075715 | 0.066
| Median |  | 0.0705 |  | 0.13628 | 0.076552 | 0.066
| Std Dev. |  | 0.0027386 |  | 0.0058172 | 0.0033132 | 0.0024495
| Rel.Std.Dev. |  | 3.885% |  | 4.236% | 4.376% | 3.711%
| PDM3 |  | -0.903% |  | 93.02% | 6.428% | -7.228%
| Z-Score (Absolute) |  | 0 | 2.65 | 8.64 | 0.67 | 0.58
| Analytical results for U in OREAS 550b (Certified Value 4.61 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 4.74 | 4.5 | 4.7 | 5.14 | 7.1934 | 4.28 | 4.83 | 4.7879 | 4.74 | 4.32 | 4.2 | 5.1 | 4.6 | 4.9684 | 4.48 | 4.34 | 5.56 | 4.4 | 4.5 | < 5 | 3.968 | 3.46 | 5.39 | 4.73 |  |  | 1
| 1 | 2 | 4.75 | 4.44 | 4.4 | 5.22 | 7.1599 | 4.45 | 5.03 | 4.8252 | 4.55 | 4.53 | 4.27 | 5.16 | 4.49 | 4.785 | 4.42 | 4.43 | 5.46 | 4.5 | 4.6 | < 5 | 4.043 | 3.42 | 5.42 | 4.61 |  |  | 13
| 1 | 3 | 4.85 | 4.88 | 4.4 | 5.26 | 7.1913 | 4.45 | 4.84 | 4.865 | 4.59 | 4.52 | 4.54 | 4.88 | 4.53 | 5.0517 | 4.6 | 4.54 | 5.59 | 4.5 | 4.5 | < 5 | 3.972 | 3.54 | 5.49 | 4.67 |  |  | 16
| 1 | 4 | 4.77 | 4.56 | 4.6 | 5.26 | 7.1558 | 4.28 | 5.04 | 4.8492 | 4.59 | 4.58 | 4.32 | 4.78 | 4.63 | 4.8718 | 4.52 | 4.52 | 5.83 | 4.5 | 4.3 | < 5 | 4.138 | 3.61 | 5.35 | 4.8 |  |  | 4.6116
| 1 | 5 | 4.74 | 4.32 | 4.4 | 5.22 | 7.1369 | 4.24 | 4.81 | 4.8338 | 4.51 | 4.24 | 4.49 | 5.12 | 4.69 | 4.7763 | 4.37 | 4.4 | 5.76 | 4.8 | 4.5 | < 5 | 4.15 | 3.82 | 5.73 | 4.75 |  |  | 199
| 1 | 6 | 4.9 | 4.16 | 4.6 | 5.24 | 7.1478 | 4.12 | 4.92 | 4.7712 | 4.52 | 4.44 | 4.38 | 4.97 | 4.48 | 4.9176 | 4.37 | 4.56 | 5.96 | 4.5 | 4.5 | < 5 | 4.061 | 3.53 | 5.58 | 4.78
| Mean |  | 4.7917 | 4.4767 | 4.5167 | 5.2233 | 7.1642 | 4.3033 | 4.9117 | 4.822 | 4.5833 | 4.4383 | 4.3667 | 5.0017 | 4.57 | 4.8951 | 4.46 | 4.465 | 5.6933 | 4.5333 | 4.4833 |  | 4.0553 | 3.5633 | 5.4933 | 4.7233
| Median |  | 4.76 | 4.47 | 4.5 | 5.23 | 7.1579 | 4.28 | 4.88 | 4.8295 | 4.57 | 4.48 | 4.35 | 5.035 | 4.565 | 4.8947 | 4.45 | 4.475 | 5.675 | 4.5 | 4.5 |  | 4.052 | 3.535 | 5.455 | 4.74
| Std Dev. |  | 0.067355 | 0.24345 | 0.13292 | 0.044572 | 0.023208 | 0.12785 | 0.10265 | 0.035983 | 0.083825 | 0.13303 | 0.13018 | 0.15079 | 0.083666 | 0.10697 | 0.090995 | 0.088034 | 0.18822 | 0.13663 | 0.098319 |  | 0.078168 | 0.14208 | 0.14152 | 0.071461
| Rel.Std.Dev. |  | 1.406% | 5.438% | 2.943% | 0.8533% | 0.324% | 2.971% | 2.09% | 0.7462% | 1.829% | 2.997% | 2.981% | 3.015% | 1.831% | 2.185% | 2.04% | 1.972% | 3.306% | 3.014% | 2.193% |  | 1.928% | 3.987% | 2.576% | 1.513%
| PDM3 |  | 3.904% | -2.926% | -2.059% | 13.26% | 55.35% | -6.685% | 6.506% | 4.563% | -0.6135% | -3.758% | -5.312% | 8.458% | -0.9026% | 6.148% | -3.288% | -3.179% | 23.46% | -1.698% | -2.782% |  | -12.06% | -22.73% | 19.12% | 2.422%
| Z-Score (Absolute) |  | 0.66 | 0.21 | 0.1 | 1.85 | 7.21 | 0.69 | 0.99 | 0.75 | 0.09 | 0.31 | 0.51 | 1.24 | 0.05 | 0.95 | 0.25 | 0.24 | 3.15 | 0.05 | 0.19 | 5.67 | 1.37 | 2.73 | 2.6 | 0.47
| Analytical results for V in OREAS 550b (Certified Value 45.9 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 1
| 1 | 1 | 50 | 42 | 58 | 57 | 61 | 45 | 39 | 48.1219 | 44 | 41 | 43 | 48 | 42 | 51.5081 | 37 | 48 | 77 | 42 | 48 | 37.6 | 46.46 | 28 | 75 | 41 |  |  | 1
| 1 | 2 | 49 | 44 | 54 | 58 | 57 | 46 | 40 | 45.654 | 44 | 43 | 42 | 49 | 40 | 48.4575 | 37.6 | 48 | 75 | 43 | 49 | 37.8 | 46.336 | 27.4 | 68 | 41 |  |  | 28
| 1 | 3 | 47 | 46 | 53 | 58 | 53 | 46 | 38 | 46.0155 | 44 | 43 | 43 | 47 | 41 | 49.9187 | 37.6 | 48 | 74 | 42 | 49 | 37.5 | 46.347 | 27.7 | 72 | 43 |  |  | 16
| 1 | 4 | 48 | 46 | 60 | 58 | 57 | 46 | 40 | 46.3037 | 44 | 43 | 41 | 48 | 40 | 50.024 | 37.7 | 49 | 76 | 45 | 48 | 37.8 | 47.765 | 29.6 | 75 | 42 |  |  | 45.9389
| 1 | 5 | 49 | 41 | 53 | 59 | 52 | 46 | 38 | 48.5596 | 44 | 42 | 41 | 47 | 41 | 48.848 | 37.3 | 48 | 70 | 42 | 49 | 37.7 | 47.941 | 29.2 | 71 | 42 |  |  | 200
| 1 | 6 | 48 | 43 | 60 | 60 | 57 | 46 | 39 | 44.6514 | 44 | 43 | 41 | 47 | 40 | 48.5704 | 37.8 | 48 | 74 | 46 | 49 | 37.5 | 46.916 | 28.8 | 74 | 42
| Mean |  | 48.5 | 43.6667 | 56.3333 | 58.3333 | 56.1667 | 45.8333 | 39 | 46.551 | 44 | 42.5 | 41.8333 | 47.6667 | 40.6667 | 49.5545 | 37.5 | 48.1667 | 74.3333 | 43.3333 | 48.6667 | 37.65 | 46.9608 | 28.45 | 72.5 | 41.8333
| Median |  | 48.5 | 43.5 | 56 | 58 | 57 | 46 | 39 | 46.1596 | 44 | 43 | 41.5 | 47.5 | 40.5 | 49.3834 | 37.6 | 48 | 74.5 | 42.5 | 49 | 37.65 | 46.688 | 28.4 | 73 | 42
| Std Dev. |  | 1.0488 | 2.0656 | 3.3862 | 1.0328 | 3.2506 | 0.40825 | 0.89443 | 1.5009 | 0 | 0.83666 | 0.98319 | 0.8165 | 0.8165 | 1.1697 | 0.29665 | 0.40825 | 2.4221 | 1.7512 | 0.5164 | 0.13784 | 0.72491 | 0.88034 | 2.7386 | 0.75277
| Rel.Std.Dev. |  | 2.162% | 4.73% | 6.011% | 1.771% | 5.787% | 0.8907% | 2.293% | 3.224% | 0% | 1.969% | 2.35% | 1.713% | 2.008% | 2.36% | 0.7911% | 0.8476% | 3.258% | 4.041% | 1.061% | 0.3661% | 1.544% | 3.094% | 3.777% | 1.799%
| PDM3 |  | 5.575% | -4.946% | 22.63% | 26.98% | 22.26% | -0.2297% | -15.1% | 1.333% | -4.221% | -7.486% | -8.937% | 3.761% | -11.48% | 7.87% | -18.37% | 4.849% | 61.81% | -5.672% | 5.938% | -18.04% | 2.225% | -38.07% | 57.82% | -8.937%
| Z-Score (Absolute) |  | 0.39 | 0.42 | 1.7 | 2.03 | 1.67 | 0.06 | 1.2 | 0.06 | 0.37 | 0.62 | 0.73 | 0.25 | 0.93 | 0.56 | 1.46 | 0.33 | 4.71 | 0.48 | 0.41 | 1.43 | 0.13 | 2.97 | 4.41 | 0.73
| Analytical results for W in OREAS 550b (Certified Value 1 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 2
| 1 | 1 | 0.98 | 0.96 | 2.1 | 0.8 | 0.81 | 1.2 | 0.98301 | 1.12 | 0.98 | 1.07 | 2.03 | 1.01 | 1.06 | 0.8 | 0.7 | 0.7 | 0.7 | < 5 | 1.025 | 2.1 | 1.14 | 1.49 |  |  |  |  | 1
| 1 | 2 | 1.23 | 0.77 | 0.9 | 1.1 | 0.83 | 1.2 | 1.1115 | 1.09 | 1.22 | 1.12 | 0.91 | 1.67 | 1.0265 | 0.9 | 0.9 | 1.2 | 0.8 | < 5 | 1.034 | 2.2 | 1.17 | 4.41 |  |  |  |  | 14
| 1 | 3 | 1.29 | 0.88 | 2.1 | 0.7 | 0.86 | 1.1 | 1.2857 | 1.25 | 1.02 | 1.31 | 0.93 | 1.03 | 1.0559 | 0.9 | 0.8 | 1 | 0.7 | < 5 | 1.004 | 2.2 | 1.19 | 1.7 |  |  |  |  | 16
| 1 | 4 | 0.91 | 0.76 | 1 | 1 | 0.85 | 1.2 | 1.2131 | 1.5 | 0.99 | 1.13 | 1.18 | 1.02 | 1.103 | 0.9 | 0.6 | 0.7 | 0.7 | < 5 | 2.13 | 2.3 | 0.99 | 1.57 |  |  |  |  | 1.0022
| 1 | 5 | 1.01 | 0.91 | 1.5 | 1 | 0.84 | 1.1 | 0.95166 | 1 | 1.54 | 1.13 | 0.94 | 1.08 | 1.1282 | 0.9 | 0.9 | 1.1 | 0.7 | < 5 | 1.078 | 2.3 | 1.13 | 1.85 |  |  |  |  | 201
| 1 | 6 | 1.13 | 0.88 | 0.7 | 1 | 0.89 | 1.2 | 0.97114 | 1.08 | 1.06 | 1.16 | 0.94 | 1.84 | 1.1369 | 0.9 | 0.7 | 0.7 | 0.8 | < 5 | 1.088 | 2.8 | 1.34 | 2.46
| Mean |  | 1.0917 | 0.86 | 1.3833 | 0.93333 | 0.84667 | 1.1667 | 1.086 | 1.1733 | 1.135 | 1.1533 | 1.155 | 1.275 | 1.0851 | 0.88333 | 0.76667 | 0.9 | 0.73333 |  | 1.2265 | 2.3167 | 1.16 | 2.2467
| Median |  | 1.07 | 0.88 | 1.25 | 1 | 0.845 | 1.2 | 1.0473 | 1.105 | 1.04 | 1.13 | 0.94 | 1.055 | 1.0815 | 0.9 | 0.75 | 0.85 | 0.7 |  | 1.056 | 2.25 | 1.155 | 1.775
| Std Dev. |  | 0.14972 | 0.079246 | 0.61455 | 0.15055 | 0.027325 | 0.05164 | 0.14039 | 0.17952 | 0.21686 | 0.082138 | 0.44031 | 0.37644 | 0.044243 | 0.040825 | 0.12111 | 0.22804 | 0.05164 |  | 0.44378 | 0.24833 | 0.11278 | 1.1146
| Rel.Std.Dev. |  | 13.71% | 9.215% | 44.43% | 16.13% | 3.227% | 4.426% | 12.93% | 15.3% | 19.11% | 7.122% | 38.12% | 29.53% | 4.077% | 4.622% | 15.8% | 25.34% | 7.042% |  | 36.18% | 10.72% | 9.723% | 49.61%
| PDM3 |  | 8.932% | -14.18% | 38.04% | -6.867% | -15.52% | 16.42% | 8.369% | 17.08% | 13.26% | 15.09% | 15.25% | 27.23% | 8.275% | -11.86% | -23.5% | -10.19% | -26.82% |  | 22.39% | 131.2% | 15.75% | 124.2%
| Z-Score (Absolute) |  | 0.21 | 1.12 | 0.94 | 0.83 | 1.17 | 0.09 | 0.23 | 0.12 | 0.04 | 0.04 | 0.04 | 0.52 | 0.23 | 1.03 | 1.49 | 0.96 | 1.62 | 5.35 | 0.32 | 4.63 | 0.06 | 4.35
| Analytical results for Y in OREAS 550b (Certified Value 4.55 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 3
| 1 | 1 | 4.64 | 4.4 | 5.9 | 6.2 | 4.57 | 4.61 | 4.947 | 4.48 | 4.26 | 4.34 | 4.84 | 4.49 | 5.2389 | 4.3 | 4.55 | 4 | < 5 | 4.499 | 3.45 | 6.29 | 4.12 | [5 empty] | 1
| 1 | 2 | 4.68 | 4.56 | 5.99 | 6.6 | 4.75 | 4.77 | 4.8624 | 4.45 | 4.4 | 4.4 | 5.04 | 4.44 | 5.0441 | 4.3 | 4.71 | 4 | < 5 | 4.495 | 3.42 | 5.96 | 4.06 | [5 empty] | 15
| 1 | 3 | 4.73 | 4.52 | 5.98 | 5.5 | 4.7 | 4.64 | 4.9231 | 4.38 | 4.46 | 4.59 | 5.11 | 4.39 | 5.1148 | 4.3 | 4.72 | 4 | < 5 | 4.462 | 3.46 | 6.24 | 4.16 | [5 empty] | 16
| 1 | 4 | 4.73 | 4.5 | 5.91 | 5.7 | 4.48 | 4.74 | 5.0503 | 4.46 | 4.37 | 4.56 | 4.77 | 4.48 | 5.1332 | 4.3 | 4.68 | 4 | < 5 | 4.621 | 3.55 | 6.48 | 3.96 | [5 empty] | 4.5461
| 1 | 5 | 4.62 | 4.34 | 5.94 | 6 | 4.88 | 4.6 | 5.0173 | 4.37 | 4.26 | 4.45 | 4.9 | 4.5 | 5.0199 | 4.4 | 4.61 | 4 | < 5 | 4.591 | 3.62 | 6.34 | 3.96 | [5 empty] | 202
| 1 | 6 | 4.64 | 4.56 | 5.89 | 5.2 | 4.63 | 4.77 | 4.8098 | 4.36 | 4.4 | 4.48 | 4.93 | 4.5 | 5.132 | 4.3 | 4.75 | 4 | < 5 | 4.474 | 3.51 | 6.37 | 3.89
| Mean |  | 4.6733 | 4.48 | 5.935 | 5.8667 | 4.6683 | 4.6883 | 4.935 | 4.4167 | 4.3583 | 4.47 | 4.9317 | 4.4667 | 5.1138 | 4.3167 | 4.67 | 4 |  | 4.5237 | 3.5017 | 6.28 | 4.025
| Median |  | 4.66 | 4.51 | 5.925 | 5.85 | 4.665 | 4.69 | 4.9351 | 4.415 | 4.385 | 4.465 | 4.915 | 4.485 | 5.1234 | 4.3 | 4.695 | 4 |  | 4.497 | 3.485 | 6.315 | 4.01
| Std Dev. |  | 0.048028 | 0.090333 | 0.042308 | 0.50465 | 0.14077 | 0.080353 | 0.09088 | 0.052409 | 0.081588 | 0.094657 | 0.12576 | 0.043665 | 0.077498 | 0.040825 | 0.075631 | 0 |  | 0.06589 | 0.07414 | 0.17652 | 0.10502
| Rel.Std.Dev. |  | 1.028% | 2.016% | 0.7129% | 8.602% | 3.015% | 1.714% | 1.842% | 1.187% | 1.872% | 2.118% | 2.55% | 0.9776% | 1.515% | 0.9457% | 1.62% | 0% |  | 1.457% | 2.117% | 2.811% | 2.609%
| PDM3 |  | 2.799% | -1.454% | 30.55% | 29.05% | 2.689% | 3.129% | 8.555% | -2.847% | -4.13% | -1.674% | 8.481% | -1.747% | 12.49% | -5.047% | 2.726% | -12.01% |  | -0.4933% | -22.97% | 38.14% | -11.46%
| Z-Score (Absolute) |  | 0.49 | 0.14 | 4.6 | 4.37 | 0.47 | 0.54 | 1.34 | 0.35 | 0.54 | 0.17 | 1.33 | 0.19 | 1.92 | 0.67 | 0.48 | 1.71 | 6.59 | 0 | 3.33 | 5.72 | 1.62
| Analytical results for Yb in OREAS 550b (Certified Value 0.43 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 05 | 08 | 15 | 16 | 17 | 22 | 23 | 24 | 26 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | [15 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.4g | [15 empty] | 3
| 1 | 1 | 0.434 | 0.5 | 0.82008 | 0.4 | 0.47492 | 0.4 | 0.4 | < 5 | 0.426 | 0.37 | 0.5 | [15 empty] | 1
| 1 | 2 | 0.465 | 0.4 | 0.8196 | 0.4 | 0.46257 | 0.44 | 0.4 | < 5 | 0.412 | 0.37 | 0.4 | [15 empty] | 16
| 1 | 3 | 0.45 | 0.4 | 0.8468 | 0.4 | 0.48437 | 0.46 | 0.4 | < 5 | 0.416 | 0.38 | 0.5 | [15 empty] | 16
| 1 | 4 | 0.446 | 0.5 | 0.86072 | 0.4 | 0.44681 | 0.43 | 0.4 | < 5 | 0.43 | 0.4 | 0.5 | [15 empty] | 0.42848
| 1 | 5 | 0.466 | 0.4 | 0.85448 | 0.4 | 0.47975 | 0.46 | 0.4 | < 5 | 0.421 | 0.39 | 0.4 | [15 empty] | 203
| 1 | 6 | 0.449 | 0.4 | 0.81176 | 0.4 | 0.47324 | 0.41 | 0.4 | < 5 | 0.411 | 0.38 | 0.5
| Mean |  | 0.45167 | 0.43333 | 0.83557 | 0.4 | 0.47028 | 0.43333 | 0.4 |  | 0.41933 | 0.38167 | 0.46667
| Median |  | 0.4495 | 0.4 | 0.83344 | 0.4 | 0.47408 | 0.435 | 0.4 |  | 0.4185 | 0.38 | 0.5
| Std Dev. |  | 0.012144 | 0.05164 | 0.020872 | 6.0809e-17 | 0.01363 | 0.025033 | 6.0809e-17 |  | 0.0076855 | 0.01169 | 0.05164
| Rel.Std.Dev. |  | 2.689% | 11.92% | 2.498% | 1.52e-14% | 2.898% | 5.777% | 1.52e-14% |  | 1.833% | 3.063% | 11.07%
| PDM3 |  | 5.413% | 1.134% | 95.01% | -6.646% | 9.756% | 1.134% | -6.646% |  | -2.134% | -10.92% | 8.913%
| Z-Score (Absolute) |  | 0.37 | 0 | 8.14 | 0.67 | 0.75 | 0 | 0.67 | 41.81 | 0.28 | 1.05 | 0.67
| Analytical results for Zn in OREAS 550b (Certified Value 22.9 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 2
| 1 | 1 | 26 | 25.9 | 21.6 | 21 | 27 | 23 | 23 | 22.4512 | 22 | 22 | 22 | 27 | 22 | 23.8102 | 23.2 | 23 | 24 | 26 | 21 | 23 | 25.7 | 26.922 | 23.7 | 21.2 | 25 |  | 1
| 1 | 2 | 22 | 24.5 | 20.8 | 21 | 28 | 23 | 21 | 20.5365 | 22 | 21 | 23 | 23 | 21 | 23.5368 | 22.7 | 22 | 24 | 25 | 21 | 23 | 25.6 | 27.34 | 23.9 | 22.4 | 25 |  | 17
| 1 | 3 | 23 | 24.8 | 22 | 22 | 24 | 24 | 22 | 21.3604 | 23 | 22 | 22 | 23 | 21 | 23.4022 | 22.6 | 22 | 24 | 24 | 21 | 23 | 26.6 | 26.235 | 23.8 | 21 | 27 |  | 16
| 1 | 4 | 23 | 27 | 21.1 | 22 | 24 | 22 | 23 | 21.7269 | 22 | 21 | 21 | 23 | 21 | 23.736 | 22.4 | 22 | 25 | 24 | 22 | 23 | 24.9 | 27.093 | 24.8 | 21.2 | 25 |  | 22.872
| 1 | 5 | 23 | 23.6 | 20.7 | 22 | 25 | 23 | 21 | 22.1994 | 22 | 21 | 22 | 23 | 21 | 22.983 | 22.7 | 22 | 25 | 23 | 22 | 25 | 24.4 | 26.544 | 24.6 | 22 | 27 |  | 204
| 1 | 6 | 22 | 21.9 | 20.1 | 22 | 22 | 23 | 22 | 20.85 | 22 | 22 | 22 | 23 | 21 | 22.9763 | 21.9 | 22 | 25 | 23 | 22 | 23 | 23.7 | 26.244 | 23.9 | 22.1 | 25
| Mean |  | 23.1667 | 24.6167 | 21.05 | 21.6667 | 25 | 23 | 22 | 21.5207 | 22.1667 | 21.5 | 22 | 23.6667 | 21.1667 | 23.4074 | 22.5833 | 22.1667 | 24.5 | 24.1667 | 21.5 | 23.3333 | 25.15 | 26.7297 | 24.1167 | 21.65 | 25.6667
| Median |  | 23 | 24.65 | 20.95 | 22 | 24.5 | 23 | 22 | 21.5436 | 22 | 21.5 | 22 | 23 | 21 | 23.4695 | 22.65 | 22 | 24.5 | 24 | 21.5 | 23 | 25.25 | 26.733 | 23.9 | 21.6 | 25
| Std Dev. |  | 1.472 | 1.777 | 0.6775 | 0.5164 | 2.1909 | 0.63246 | 0.89443 | 0.75009 | 0.40825 | 0.54772 | 0.63246 | 1.633 | 0.40825 | 0.36136 | 0.42622 | 0.40825 | 0.54772 | 1.169 | 0.54772 | 0.8165 | 1.033 | 0.45968 | 0.46224 | 0.58566 | 1.0328
| Rel.Std.Dev. |  | 6.354% | 7.219% | 3.219% | 2.383% | 8.764% | 2.75% | 4.066% | 3.485% | 1.842% | 2.548% | 2.875% | 6.9% | 1.929% | 1.544% | 1.887% | 1.842% | 2.236% | 4.837% | 2.548% | 3.499% | 4.107% | 1.72% | 1.917% | 2.705% | 4.024%
| PDM3 |  | 1.288% | 7.628% | -7.966% | -5.27% | 9.304% | 0.5596% | -3.813% | -5.908% | -3.084% | -5.999% | -3.813% | 3.474% | -7.456% | 2.341% | -1.262% | -3.084% | 7.118% | 5.66% | -5.999% | 2.017% | 9.96% | 16.87% | 5.442% | -5.343% | 12.22%
| Z-Score (Absolute) |  | 0.08 | 0.82 | 0.99 | 0.67 | 1.01 | 0 | 0.51 | 0.75 | 0.42 | 0.76 | 0.51 | 0.34 | 0.93 | 0.21 | 0.21 | 0.42 | 0.76 | 0.59 | 0.76 | 0.17 | 1.09 | 1.89 | 0.56 | 0.68 | 1.35
| Analytical results for Zr in OREAS 550b (Certified Value 11 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 1
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 1
| 1 | 1 | 11.1 | 6.9 | 9.8 | 14.5 | 11 | 11.6 | 10.2944 | 9.5 | 10.3 | 10.3 | 10.9 | 10.3 | 12.1642 | 9.5 | 11.2 | 9.3 | 13.5 | 12.016 | 6.91 | 13.7 | 10.1 | [5 empty] | 1
| 1 | 2 | 10.9 | 10.1 | 10.4 | 15.4 | 11 | 11.8 | 10.6646 | 9.6 | 10.6 | 10.6 | 11.1 | 10.2 | 11.8738 | 9.6 | 11.2 | 9.3 | 13.8 | 12.039 | 7.17 | 11.1 | 10.1 | [5 empty] | 22
| 1 | 3 | 10.9 | 9.6 | 10.7 | 13 | 11 | 11.1 | 10.3206 | 9.5 | 10.6 | 10.8 | 11.2 | 10.5 | 12.1771 | 9.9 | 11.5 | 9.3 | 13.8 | 11.99 | 7.27 | 12.8 | 10.4 | [5 empty] | 16
| 1 | 4 | 11.1 | 8.7 | 10.7 | 13.3 | 11 | 11.6 | 10.453 | 9.6 | 10.5 | 10.4 | 10.8 | 10.7 | 12.1333 | 9.9 | 11.3 | 9.3 | 13.5 | 12.575 | 7.3 | 13.3 | 9.9 | [5 empty] | 10.9703
| 1 | 5 | 10.9 | 8.5 | 10.8 | 14.3 | 11 | 11.2 | 10.7725 | 9.4 | 10.4 | 10.6 | 10.9 | 10.1 | 11.9968 | 9.5 | 11.1 | 9.4 | 13.2 | 12.202 | 7.35 | 10.3 | 10 | [5 empty] | 205
| 1 | 6 | 10.7 | 3.1 | 10.5 | 12.4 | 11 | 11.8 | 10.4915 | 9.5 | 10.6 | 10.7 | 11.1 | 10.6 | 12.0959 | 9.6 | 11.3 | 9.3 | 13.8 | 12.213 | 7.04 | 13.7 | 10.4
| Mean |  | 10.9333 | 7.8167 | 10.4833 | 13.8167 | 11 | 11.5167 | 10.4994 | 9.5167 | 10.5 | 10.5667 | 11 | 10.4 | 12.0735 | 9.6667 | 11.2667 | 9.3167 | 13.6 | 12.1725 | 7.1733 | 12.4833 | 10.15
| Median |  | 10.9 | 8.6 | 10.6 | 13.8 | 11 | 11.6 | 10.4723 | 9.5 | 10.55 | 10.6 | 11 | 10.4 | 12.1146 | 9.6 | 11.25 | 9.3 | 13.65 | 12.1205 | 7.22 | 13.05 | 10.1
| Std Dev. |  | 0.15055 | 2.5585 | 0.3656 | 1.1089 | 0 | 0.29944 | 0.18874 | 0.075277 | 0.12649 | 0.18619 | 0.15492 | 0.23664 | 0.11729 | 0.18619 | 0.13663 | 0.040825 | 0.24495 | 0.21915 | 0.16955 | 1.4428 | 0.20736
| Rel.Std.Dev. |  | 1.377% | 32.73% | 3.487% | 8.026% | 0% | 2.6% | 1.798% | 0.791% | 1.205% | 1.762% | 1.408% | 2.275% | 0.9715% | 1.926% | 1.213% | 0.4382% | 1.801% | 1.8% | 2.364% | 11.56% | 2.043%
| PDM3 |  | -0.3367% | -28.75% | -4.439% | 25.95% | 0.271% | 4.981% | -4.292% | -13.25% | -4.287% | -3.679% | 0.271% | -5.198% | 10.06% | -11.88% | 2.702% | -15.07% | 23.97% | 10.96% | -34.61% | 13.79% | -7.477%
| Z-Score (Absolute) |  | 0.27 | 2.06 | 0.06 | 2.44 | 0.32 | 0.71 | 0.05 | 0.79 | 0.05 | 0 | 0.32 | 0.12 | 1.13 | 0.67 | 0.52 | 0.94 | 2.27 | 1.2 | 2.54 | 1.44 | 0.31
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C649: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C667: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C703: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C721: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C739: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C757: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C775: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C793: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C811: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C829: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C847: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C865: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C883: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C901: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C919: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C937: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C955: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C973: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C991: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1171: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for As in OREAS 550b (Indicative Value < 10 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 5
| 1 | 3 | < 10 | 16
| 1 | 4 | < 10 | < 10
| 1 | 5 | < 10 | 34
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 550b (Indicative Value 0.135 wt.%) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 24 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.142 | 1
| 1 | 2 | 0.127 | 29
| 1 | 3 | 0.13 | 16
| 1 | 4 | 0.14 | 0.13517
| 1 | 5 | 0.13 | 35
| 1 | 6 | 0.142
| Mean |  | 0.13517
| Median |  | 0.135
| Std Dev. |  | 0.0068823
| Rel.Std.Dev. |  | 5.092%
| PDM3 |  | -2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 550b (Indicative Value 1.81 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 24 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.83 | 1
| 1 | 2 | 1.82 | 30
| 1 | 3 | 1.8 | 16
| 1 | 4 | 1.83 | 1.8133
| 1 | 5 | 1.8 | 36
| 1 | 6 | 1.8
| Mean |  | 1.8133
| Median |  | 1.81
| Std Dev. |  | 0.015055
| Rel.Std.Dev. |  | 0.8303%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 550b (Indicative Value 73 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 80 | 28
| 1 | 3 | 70 | 16
| 1 | 4 | < 10 | 72.5
| 1 | 5 | 120 | 34
| 1 | 6 | 60
| Mean |  | 86
| Median |  | 80
| Std Dev. |  | 24.0832
| Rel.Std.Dev. |  | 28%
| PDM3 |  | 18.62%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 550b (Indicative Value 53 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 50 | 13
| 1 | 3 | 70 | 16
| 1 | 4 | 110 | 52.5
| 1 | 5 | 10 | 35
| 1 | 6 | < 10
| Mean |  | 62
| Median |  | 70
| Std Dev. |  | 36.3318
| Rel.Std.Dev. |  | 58.6%
| PDM3 |  | 18.1%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 550b (Indicative Value 1.16 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 24 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.16 | 1
| 1 | 2 | 1.13 | 30
| 1 | 3 | 1.18 | 16
| 1 | 4 | 1.18 | 1.1567
| 1 | 5 | 1.15 | 36
| 1 | 6 | 1.14
| Mean |  | 1.1567
| Median |  | 1.155
| Std Dev. |  | 0.020656
| Rel.Std.Dev. |  | 1.786%
| PDM3 |  | -3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 550b (Indicative Value 72 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 110 | 1
| 1 | 2 | 70 | 28
| 1 | 3 | 30 | 16
| 1 | 4 | 70 | 71.6667
| 1 | 5 | 90 | 34
| 1 | 6 | 60
| Mean |  | 71.6667
| Median |  | 70
| Std Dev. |  | 27.1416
| Rel.Std.Dev. |  | 37.87%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 550b (Indicative Value 75 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 70 | 17
| 1 | 3 | 90 | 16
| 1 | 4 | 80 | 75
| 1 | 5 | 70 | 35
| 1 | 6 | 70
| Mean |  | 75
| Median |  | 70
| Std Dev. |  | 8.3666
| Rel.Std.Dev. |  | 11.16%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.