## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 553b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 0.10281 | 0.084778 | 0.12083 | IND | IND
Al, Aluminium (wt.%) | 6.4086 | 6.2536 | 6.5637 | 6.2877 | 6.5295
As, Arsenic (ppm) | 53.5899 | 50.5908 | 56.589 | 52.0117 | 55.1681
Ba, Barium (ppm) | 261.8487 | 251.0808 | 272.6166 | 255.3024 | 268.3949
Be, Beryllium (ppm) | 4.7054 | 4.5552 | 4.8557 | 4.5759 | 4.835
Bi, Bismuth (ppm) | 1.6997 | 1.6196 | 1.7798 | 1.6507 | 1.7487
Ca, Calcium (wt.%) | 0.31286 | 0.30386 | 0.32185 | 0.30535 | 0.32037
Cd, Cadmium (ppm) | 0.29591 | 0.26969 | 0.32212 | 0.27058 | 0.32123
Ce, Cerium (ppm) | 144.2478 | 134.951 | 153.5446 | 139.5394 | 148.9562
Co, Cobalt (wt.%) | 0.58712 | 0.57679 | 0.59746 | 0.58088 | 0.59336
Cr, Chromium (ppm) | 90.5935 | 85.0407 | 96.1463 | 87.645 | 93.542
Cs, Caesium (ppm) | 2.0815 | 1.9806 | 2.1824 | 1.9999 | 2.1631
Cu, Copper (wt.%) | 3.6027 | 3.5421 | 3.6634 | 3.569 | 3.6365
Dy, Dysprosium (ppm) | 4.4419 | 3.9757 | 4.9082 | 4.2397 | 4.6442
Er, Erbium (ppm) | 2.8486 | 2.5375 | 3.1597 | 2.7337 | 2.9636
Eu, Europium (ppm) | 0.80926 | 0.67589 | 0.94264 | 0.77245 | 0.84608
Fe, Iron (wt.%) | 3.1829 | 3.1155 | 3.2504 | 3.1338 | 3.2321
Ga, Gallium (ppm) | 27.9047 | 26.736 | 29.0733 | 27.2261 | 28.5832
Gd, Gadolinium (ppm) | 5.006 | 4.3124 | 5.6997 | 4.8625 | 5.1495
Hf, Hafnium (ppm) | 5.4611 | 5.1949 | 5.7274 | 5.3026 | 5.6197
Ho, Holmium (ppm) | 0.96314 | 0.81788 | 1.1084 | 0.92671 | 0.99958
In, Indium (ppm) | 2.1391 | 2.0646 | 2.2135 | 2.0751 | 2.2031
K, Potassium (wt.%) | 2.5541 | 2.4941 | 2.6142 | 2.4988 | 2.6095
La, Lanthanum (ppm) | 68.2929 | 62.9139 | 73.6719 | 66.1585 | 70.4273
Li, Lithium (ppm) | 62.6357 | 60.9162 | 64.3551 | 61.1807 | 64.0906
Lu, Lutetium (ppm) | 0.41808 | 0.35902 | 0.47714 | 0.39603 | 0.44013
Mg, Magnesium (wt.%) | 3.5765 | 3.4859 | 3.6671 | 3.4995 | 3.6535
Mn, Manganese (wt.%) | 0.013474 | 0.013199 | 0.013748 | 0.013136 | 0.013811
Mo, Molybdenum (ppm) | 15.0823 | 14.4989 | 15.6657 | 14.7176 | 15.447
Na, Sodium (wt.%) | 0.044918 | 0.042089 | 0.047747 | 0.043139 | 0.046696
Nb, Niobium (ppm) | 17.3826 | 15.3833 | 19.382 | 16.586 | 18.1793
Nd, Neodymium (ppm) | 62.7385 | 55.9279 | 69.5491 | 61.3764 | 64.1007
Ni, Nickel (ppm) | 124.4568 | 121.7332 | 127.1805 | 122.7978 | 126.1158
P, Phosphorus (wt.%) | 0.03505 | 0.033809 | 0.036291 | 0.034051 | 0.036049
Pb, Lead (ppm) | 11.234 | 10.5022 | 11.9658 | 10.7522 | 11.7157
Pr, Praseodymium (ppm) | 16.8542 | 15.2039 | 18.5044 | 16.4289 | 17.2795
Rb, Rubidium (ppm) | 95.8242 | 90.3221 | 101.3262 | 92.4138 | 99.2345
Re, Rhenium (ppm) | 0.027479 | 0.023559 | 0.0314 | 0.023724 | 0.031235
S, Sulphur (wt.%) | 3.3376 | 3.273 | 3.4021 | 3.2892 | 3.386
Sb, Antimony (ppm) | 3.0122 | 2.7864 | 3.238 | 2.8428 | 3.1816
Sc, Scandium (ppm) | 12.5054 | 11.8561 | 13.1547 | 12.1928 | 12.818
Se, Selenium (ppm) | 5.729 | 4.9556 | 6.5023 | 5.1613 | 6.2966
Sm, Samarium (ppm) | 8.6857 | 7.766 | 9.6055 | 8.416 | 8.9555
Sn, Tin (ppm) | 2.8123 | 2.6636 | 2.961 | 2.6518 | 2.9728
Sr, Strontium (ppm) | 70.0375 | 67.1551 | 72.92 | 68.1853 | 71.8898
Ta, Tantalum (ppm) | 1.2644 | 1.0947 | 1.4341 | 1.2 | 1.3288
Tb, Terbium (ppm) | 0.74168 | 0.65509 | 0.82826 | 0.68943 | 0.79393
Te, Tellurium (ppm) | 0.063233 | 0.048903 | 0.077563 | IND | IND
Th, Thorium (ppm) | 15.4502 | 14.8224 | 16.0779 | 14.8227 | 16.0776
Ti, Titanium (wt.%) | 0.25811 | 0.22295 | 0.29327 | 0.24974 | 0.26648
Tl, Thallium (ppm) | 0.66915 | 0.63353 | 0.70477 | 0.64249 | 0.69581
Tm, Thulium (ppm) | 0.4214 | 0.32967 | 0.51313 | 0.38667 | 0.45614
U, Uranium (ppm) | 8.6004 | 8.2558 | 8.9451 | 8.3015 | 8.8994
V, Vanadium (ppm) | 348.4629 | 339.2868 | 357.639 | 341.0423 | 355.8835
W, Tungsten (ppm) | 2.52 | 2.3219 | 2.718 | 2.3666 | 2.6733
Y, Yttrium (ppm) | 23.7694 | 22.4891 | 25.0497 | 23.0842 | 24.4546
Yb, Ytterbium (ppm) | 2.8396 | 2.5305 | 3.1486 | 2.734 | 2.9451
Zn, Zinc (ppm) | 16.1493 | 15.1005 | 17.1982 | 15.4755 | 16.8232
Zr, Zirconium (ppm) | 187.8487 | 181.1021 | 194.5954 | 183.4002 | 192.2973
Sulphuric Acid 5% Leach
Co, Cobalt (wt.%) | 0.12698 | 0.11894 | 0.13501 | 0.12372 | 0.13023
Cu, Copper (wt.%) | 1.0662 | 1.0199 | 1.1126 | 1.0402 | 1.0923
Infrared Combustion
C, Carbon (wt.%) | 4.5091 | 4.4295 | 4.5887 | 4.4689 | 4.5493
S, Sulphur (wt.%) | 3.402 | 3.3499 | 3.4541 | 3.3682 | 3.4358
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 12.5052 | 12.3171 | 12.6932 | 12.3715 | 12.6388
BaO, Barium oxide (ppm) | 303.4666 | 241.4474 | 365.4859 | IND | IND
CaO, Calcium oxide (wt.%) | 0.43499 | 0.42342 | 0.44656 | 0.42066 | 0.44933
Co, Cobalt (wt.%) | 0.58666 | 0.57219 | 0.60112 | 0.57671 | 0.5966
Cu, Copper (wt.%) | 3.6212 | 3.555 | 3.6873 | 3.5699 | 3.6724
Fe2O3, Iron(III) oxide (wt.%) | 4.5996 | 4.5339 | 4.6653 | 4.548 | 4.6512
K2O, Potassium oxide (wt.%) | 3.043 | 2.9885 | 3.0975 | 3.0091 | 3.0769
MgO, Magnesium oxide (wt.%) | 5.9391 | 5.8372 | 6.041 | 5.8629 | 6.0153
MnO, Manganese oxide (wt.%) | 0.016648 | 0.014181 | 0.019114 | IND | IND
Ni, Nickel (ppm) | 125.2575 | 86.88 | 163.635 | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.082926 | 0.073547 | 0.092305 | 0.079177 | 0.086675
SiO2, Silicon dioxide (wt.%) | 53.7231 | 53.166 | 54.2801 | 53.2835 | 54.1626
SO3, Sulphur trioxide (wt.%) | 8.5942 | 8.3707 | 8.8177 | 8.4408 | 8.7475
SrO, Strontium oxide (ppm) | 95.6006 | 75.025 | 116.1761 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.79148 | 0.77479 | 0.80818 | 0.7742 | 0.80877
V2O5, Vanadium(V) oxide (ppm) | 614.1301 | 528.4007 | 699.8595 | IND | IND
Zr, Zirconium (ppm) | 240.4354 | 186.6997 | 294.1711 | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 12.7221 | 12.5286 | 12.9155 | 12.6365 | 12.8076
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.4666 | 6.3074 | 6.6258 | 6.3612 | 6.572
As, Arsenic (ppm) | 52.866 | 46.6421 | 59.0899 | 50.1636 | 55.5684
B, Boron (ppm) | 217.1039 | 196.204 | 238.0038 | 210.061 | 224.1468
Ba, Barium (ppm) | 264.0065 | 253.7507 | 274.2622 | 256.7242 | 271.2888
Be, Beryllium (ppm) | 5.4194 | 4.6172 | 6.2217 | IND | IND
Bi, Bismuth (ppm) | 1.7722 | 1.6005 | 1.944 | IND | IND
Ca, Calcium (wt.%) | 0.3067 | 0.27969 | 0.33371 | 0.29245 | 0.32094
Ce, Cerium (ppm) | 151.6062 | 144.2255 | 158.9869 | 147.0153 | 156.1971
Co, Cobalt (wt.%) | 0.58122 | 0.56536 | 0.59708 | 0.5739 | 0.58854
Cr, Chromium (ppm) | 110.467 | 97.4207 | 123.5133 | 102.3418 | 118.5922
Cs, Caesium (ppm) | 2.1337 | 2.014 | 2.2534 | 2.0234 | 2.244
Cu, Copper (wt.%) | 3.6061 | 3.5006 | 3.7116 | 3.5494 | 3.6627
Dy, Dysprosium (ppm) | 7.0307 | 6.6113 | 7.4501 | 6.6588 | 7.4026
Er, Erbium (ppm) | 3.7332 | 3.3901 | 4.0764 | 3.498 | 3.9684
Eu, Europium (ppm) | 0.91026 | 0.81751 | 1.003 | 0.8193 | 1.0012
Fe, Iron (wt.%) | 3.1677 | 3.0841 | 3.2514 | 3.1177 | 3.2178
Ga, Gallium (ppm) | 27.6572 | 26.3118 | 29.0027 | 26.3953 | 28.9191
Gd, Gadolinium (ppm) | 6.8708 | 6.4495 | 7.2922 | 6.4352 | 7.3065
Ge, Germanium (ppm) | 2.2382 | 1.3532 | 3.1231 | IND | IND
Ho, Holmium (ppm) | 1.352 | 1.2095 | 1.4945 | 1.1766 | 1.5274
In, Indium (ppm) | 2.2177 | 2.0331 | 2.4023 | 2.0651 | 2.3703
K, Potassium (wt.%) | 2.5832 | 2.5131 | 2.6534 | 2.516 | 2.6505
La, Lanthanum (ppm) | 76.9894 | 73.4214 | 80.5574 | 74.2385 | 79.7403
Li, Lithium (ppm) | 64.8582 | 60.4771 | 69.2393 | 63.0137 | 66.7027
Lu, Lutetium (ppm) | 0.55961 | 0.46926 | 0.64995 | 0.50447 | 0.61475
Mg, Magnesium (wt.%) | 3.5675 | 3.4824 | 3.6526 | 3.475 | 3.66
Mn, Manganese (wt.%) | 0.014047 | 0.013232 | 0.014863 | 0.013679 | 0.014415
Mo, Molybdenum (ppm) | 15.9784 | 14.5219 | 17.4349 | IND | IND
Nb, Niobium (ppm) | 25.5676 | 23.8188 | 27.3164 | 23.9626 | 27.1725
Nd, Neodymium (ppm) | 64.6418 | 62.1097 | 67.1739 | 62.1217 | 67.1619
Ni, Nickel (ppm) | 125.636 | 115.6925 | 135.5796 | 118.8917 | 132.3803
P, Phosphorus (wt.%) | 0.037287 | 0.026972 | 0.047601 | IND | IND
Pb, Lead (ppm) | 32.8276 | 21.8531 | 43.802 | IND | IND
Pr, Praseodymium (ppm) | 18.131 | 17.2712 | 18.9908 | 17.4161 | 18.8459
Rb, Rubidium (ppm) | 96.2434 | 92.6435 | 99.8432 | 93.1823 | 99.3045
Re, Rhenium (ppm) | < 0.1 | IND | IND | IND | IND
S, Sulphur (wt.%) | 3.2919 | 3.1787 | 3.4051 | 3.2241 | 3.3597
Sb, Antimony (ppm) | 3.3495 | 2.9268 | 3.7723 | 2.9939 | 3.7052
Sc, Scandium (ppm) | 12.2481 | 11.2231 | 13.2732 | IND | IND
Si, Silicon (wt.%) | 25.1713 | 24.5671 | 25.7756 | 24.7404 | 25.6023
Sm, Samarium (ppm) | 8.9016 | 8.2134 | 9.5898 | 8.5226 | 9.2806
Sr, Strontium (ppm) | 69.6557 | 65.7709 | 73.5406 | 66.9182 | 72.3932
Ta, Tantalum (ppm) | 2.0098 | 1.6557 | 2.364 | 1.7741 | 2.2456
Tb, Terbium (ppm) | 1.1832 | 1.1008 | 1.2656 | 1.0974 | 1.2689
Th, Thorium (ppm) | 15.9533 | 15.2987 | 16.6079 | 15.3775 | 16.5292
Ti, Titanium (wt.%) | 0.46085 | 0.44664 | 0.47507 | 0.44793 | 0.47378
Tl, Thallium (ppm) | 0.7292 | 0.6137 | 0.8447 | IND | IND
Tm, Thulium (ppm) | 0.57021 | 0.51404 | 0.62639 | 0.51921 | 0.62122
U, Uranium (ppm) | 8.6995 | 8.2824 | 9.1167 | 8.4211 | 8.978
V, Vanadium (ppm) | 364.947 | 351.5757 | 378.3182 | 358.4188 | 371.4752
W, Tungsten (ppm) | 3.1755 | 2.4955 | 3.8555 | IND | IND
Y, Yttrium (ppm) | 35.6106 | 34.0966 | 37.1247 | 33.797 | 37.4243
Yb, Ytterbium (ppm) | 3.6337 | 3.2646 | 4.0029 | 3.426 | 3.8414
Aqua Regia Digestion
Ag, Silver (ppm) | 0.076744 | 0.064474 | 0.089013 | IND | IND
Al, Aluminium (wt.%) | 0.93966 | 0.86944 | 1.0099 | 0.90333 | 0.97599
As, Arsenic (ppm) | 51.3639 | 49.3509 | 53.3768 | 49.7035 | 53.0242
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 49.8917 | 46.5066 | 53.2768 | 47.1112 | 52.6722
Be, Beryllium (ppm) | 2.1563 | 2.0122 | 2.3003 | 2.0913 | 2.2213
Bi, Bismuth (ppm) | 1.5791 | 1.5129 | 1.6453 | 1.536 | 1.6223
Ca, Calcium (wt.%) | 0.30016 | 0.28947 | 0.31085 | 0.29201 | 0.30831
Cd, Cadmium (ppm) | 0.29845 | 0.27487 | 0.32202 | 0.27774 | 0.31915
Ce, Cerium (ppm) | 42.8634 | 39.417 | 46.3097 | 41.2042 | 44.5226
Co, Cobalt (wt.%) | 0.54777 | 0.52128 | 0.57426 | 0.5353 | 0.56024
Cr, Chromium (ppm) | 33.9115 | 32.0662 | 35.7568 | 32.7236 | 35.0995
Cs, Caesium (ppm) | 0.29441 | 0.24894 | 0.33988 | 0.28141 | 0.30741
Cu, Copper (wt.%) | 3.6039 | 3.5554 | 3.6525 | 3.5717 | 3.6361
Dy, Dysprosium (ppm) | 0.79906 | 0.67966 | 0.91846 | 0.77699 | 0.82113
Er, Erbium (ppm) | 0.36324 | 0.32901 | 0.39747 | 0.3433 | 0.38317
Eu, Europium (ppm) | 0.31443 | 0.26507 | 0.36378 | 0.29501 | 0.33384
Fe, Iron (wt.%) | 2.9675 | 2.889 | 3.0461 | 2.9148 | 3.0203
Ga, Gallium (ppm) | 7.6749 | 7.0081 | 8.3418 | 7.3458 | 8.004
Gd, Gadolinium (ppm) | 1.7358 | 1.4767 | 1.9949 | 1.6831 | 1.7885
Hf, Hafnium (ppm) | 0.30042 | 0.28363 | 0.31721 | 0.28229 | 0.31854
Hg, Mercury (ppm) | 0.050371 | 0.037129 | 0.063613 | IND | IND
Ho, Holmium (ppm) | 0.1392 | 0.11794 | 0.16046 | 0.13412 | 0.14428
In, Indium (ppm) | 1.5437 | 1.4612 | 1.6262 | 1.4648 | 1.6226
K, Potassium (wt.%) | 0.19011 | 0.18003 | 0.20018 | 0.18036 | 0.19985
La, Lanthanum (ppm) | 18.9644 | 17.5314 | 20.3973 | 18.2595 | 19.6692
Li, Lithium (ppm) | 35.4644 | 32.1583 | 38.7705 | 34.158 | 36.7708
Lu, Lutetium (ppm) | 0.041033 | 0.037232 | 0.044835 | 0.036458 | 0.045608
Mg, Magnesium (wt.%) | 2.8486 | 2.7283 | 2.9688 | 2.7829 | 2.9142
Mn, Manganese (wt.%) | 0.012506 | 0.012077 | 0.012935 | 0.012162 | 0.01285
Mo, Molybdenum (ppm) | 14.1916 | 13.6748 | 14.7084 | 13.8488 | 14.5345
Na, Sodium (wt.%) | 0.0097809 | 0.0089136 | 0.010648 | IND | IND
Nd, Neodymium (ppm) | 20.1566 | 15.63 | 24.6833 | 19.6132 | 20.7001
Ni, Nickel (ppm) | 115.4174 | 111.347 | 119.4877 | 112.7627 | 118.072
P, Phosphorus (wt.%) | 0.020547 | 0.01962 | 0.021474 | 0.019598 | 0.021495
Pb, Lead (ppm) | 8.42 | 7.9887 | 8.8514 | 8.0951 | 8.745
Pr, Praseodymium (ppm) | 5.2784 | 4.2189 | 6.338 | 5.1081 | 5.4488
Rb, Rubidium (ppm) | 6.2944 | 5.6579 | 6.9308 | 5.9579 | 6.6308
Re, Rhenium (ppm) | 0.025986 | 0.022847 | 0.029126 | 0.023855 | 0.028118
S, Sulphur (wt.%) | 3.2631 | 3.1773 | 3.3488 | 3.2097 | 3.3165
Sb, Antimony (ppm) | 1.8554 | 1.7136 | 1.9972 | 1.7619 | 1.9489
Sc, Scandium (ppm) | 3.2999 | 3.0786 | 3.5212 | 3.0779 | 3.5219
Se, Selenium (ppm) | 5.6527 | 5.1847 | 6.1208 | 5.215 | 6.0904
Sm, Samarium (ppm) | 3.197 | 2.6658 | 3.7281 | 3.0763 | 3.3176
Sn, Tin (ppm) | 1.2775 | 1.1879 | 1.367 | 1.2065 | 1.3485
Sr, Strontium (ppm) | 16.4063 | 15.0558 | 17.7568 | 15.6747 | 17.1378
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.19025 | 0.16984 | 0.21067 | 0.17939 | 0.20112
Te, Tellurium (ppm) | 0.047667 | 0.033357 | 0.061977 | IND | IND
Th, Thorium (ppm) | 6.1773 | 5.8007 | 6.5539 | 5.9667 | 6.3879
U, Uranium (ppm) | 3.0052 | 2.8615 | 3.1489 | 2.9198 | 3.0905
V, Vanadium (ppm) | 55.7682 | 51.2658 | 60.2705 | 53.3961 | 58.1403
W, Tungsten (ppm) | 0.96473 | 0.88059 | 1.0489 | 0.90634 | 1.0231
Y, Yttrium (ppm) | 3.0963 | 2.8761 | 3.3166 | 2.9606 | 3.2321
Yb, Ytterbium (ppm) | 0.28194 | 0.22297 | 0.3409 | 0.25723 | 0.30664
Zn, Zinc (ppm) | 13.0011 | 11.9569 | 14.0454 | 11.7975 | 14.2048
Zr, Zirconium (ppm) | 9.3918 | 8.6305 | 10.1531 | 8.9695 | 9.8141
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 553b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 20 | Hg | ppm | 0.058333
Ge | ppm | 0.26989 | Pt | ppb | 7.3333
Borate Fusion XRF
Ag | ppm | 0.0315 | Hf | ppm | 42.4732 | Se | ppm | 3.8333
As | ppm | 97.1667 | Ho | ppm | 1.375 | Sm | ppm | 8.75
Be | ppm | 5 | In | ppm | 2.2 | Sn | ppm | 32.95
Bi | ppm | < 100 | La | ppm | 81.4472 | Ta | ppm | 69.1952
Cd | ppm | < 10 | Lu | ppm | 0.545 | Tb | ppm | 1.2
Ce | ppm | 151.3228 | Mo | ppm | 24.375 | Te | ppm | 0.075
Cl | ppm | 413.3333 | Na2O | wt.% | 0.074048 | Th | ppm | 15.68
Cr2O3 | ppm | 143.353 | Nb | ppm | 44.5787 | Tl | ppm | 0.57333
Cs | ppm | 2.05 | Nd | ppm | 63.2992 | Tm | ppm | 0.55333
Dy | ppm | 6.81 | Pb | ppm | < 50 | U | ppm | 46.7366
Er | ppm | 4.0833 | Pr | ppm | 18.0183 | W | ppm | 4.8487
Eu | ppm | 0.93667 | Rb | ppm | 66.8556 | Y | ppm | 48.9983
F | ppm | < 5000 | Re | ppm | < 0.1 | Yb | ppm | 3.75
Ga | ppm | 27.2667 | Sb | ppm | 30.3687 | Zn | ppm | < 50
Gd | ppm | 7.01 | Sc | ppm | 14.6667
Peroxide Fusion ICP
Ag | ppm | 1.5608 | Na | wt.% | 0.066766 | Zn | ppm | 27.2003
Cd | ppm | 0.325 | Se | ppm | 9.75 | Zr | ppm | 241.1063
Hf | ppm | 7.9715 | Sn | ppm | 2.8557
Hg | ppm | < 5 | Te | ppm | < 1
Aqua Regia Digestion
Ge | ppm | 0.11772 | Pt | ppb | 3 | Tm | ppm | 0.059369
Nb | ppm | 0.086267 | Ti | wt.% | 0.0014163
Pd | ppb | < 10 | Tl | ppm | 0.23012
3-Acid Digestion (no HF)
As | ppm | < 10 | Ni | ppm | 101.6667 | V | ppm | 101.6667
Co | wt.% | 0.56417 | Pb | ppm | 156.6667 | Zn | ppm | 103.3333
Cu | wt.% | 3.5967 | S | wt.% | 3.32
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 553b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 0.10281 | 0.012263 | 0.07828 | 0.12733 | 0.066017 | 0.1396 | 11.93% | 23.86% | 35.78% | 0.097666 | 0.10795
Al, wt.% | 6.4086 | 0.20522 | 5.9982 | 6.8191 | 5.7929 | 7.0243 | 3.202% | 6.405% | 9.607% | 6.0882 | 6.7291
As, ppm | 53.5899 | 2.9146 | 47.7608 | 59.4191 | 44.8462 | 62.3336 | 5.439% | 10.88% | 16.32% | 50.9104 | 56.2694
Ba, ppm | 261.8487 | 13.5406 | 234.7675 | 288.9298 | 221.2269 | 302.4704 | 5.171% | 10.34% | 15.51% | 248.7562 | 274.9411
Be, ppm | 4.7054 | 0.1873 | 4.3308 | 5.08 | 4.1435 | 5.2673 | 3.981% | 7.961% | 11.94% | 4.4701 | 4.9407
Bi, ppm | 1.6997 | 0.080514 | 1.5387 | 1.8607 | 1.4581 | 1.9412 | 4.737% | 9.474% | 14.21% | 1.6147 | 1.7847
Ca, wt.% | 0.31286 | 0.010765 | 0.29133 | 0.33439 | 0.28056 | 0.34515 | 3.441% | 6.882% | 10.32% | 0.29721 | 0.3285
Cd, ppm | 0.29591 | 0.016015 | 0.26388 | 0.32794 | 0.24786 | 0.34395 | 5.412% | 10.82% | 16.24% | 0.28111 | 0.3107
Ce, ppm | 144.2478 | 14.5831 | 115.0816 | 173.414 | 100.4985 | 187.9971 | 10.11% | 20.22% | 30.33% | 137.0354 | 151.4602
Co, wt.% | 0.58712 | 0.012256 | 0.56261 | 0.61163 | 0.55036 | 0.62389 | 2.087% | 4.175% | 6.262% | 0.55777 | 0.61648
Cr, ppm | 90.5935 | 10.323 | 69.9476 | 111.2394 | 59.6246 | 121.5624 | 11.39% | 22.79% | 34.18% | 86.0638 | 95.1232
Cs, ppm | 2.0815 | 0.10587 | 1.8698 | 2.2932 | 1.7639 | 2.3991 | 5.086% | 10.17% | 15.26% | 1.9774 | 2.1856
Cu, wt.% | 3.6027 | 0.06166 | 3.4794 | 3.7261 | 3.4178 | 3.7877 | 1.711% | 3.423% | 5.134% | 3.4226 | 3.7829
Dy, ppm | 4.4419 | 0.318 | 3.8059 | 5.0779 | 3.4879 | 5.3959 | 7.159% | 14.32% | 21.48% | 4.2198 | 4.664
Er, ppm | 2.8486 | 0.28103 | 2.2866 | 3.4107 | 2.0055 | 3.6917 | 9.866% | 19.73% | 29.6% | 2.7062 | 2.9911
Eu, ppm | 0.80926 | 0.11159 | 0.58609 | 1.0324 | 0.47451 | 1.144 | 13.79% | 27.58% | 41.37% | 0.7688 | 0.84973
Fe, wt.% | 3.1829 | 0.077559 | 3.0278 | 3.338 | 2.9503 | 3.4156 | 2.437% | 4.873% | 7.31% | 3.0238 | 3.3421
Ga, ppm | 27.9047 | 1.241 | 25.4226 | 30.3868 | 24.1815 | 31.6278 | 4.447% | 8.895% | 13.34% | 26.5094 | 29.2999
Gd, ppm | 5.006 | 0.62818 | 3.7497 | 6.2624 | 3.1215 | 6.8906 | 12.55% | 25.1% | 37.65% | 4.7557 | 5.2563
Hf, ppm | 5.4611 | 0.21009 | 5.0409 | 5.8813 | 4.8308 | 6.0914 | 3.847% | 7.694% | 11.54% | 5.1881 | 5.7342
Ho, ppm | 0.96314 | 0.097831 | 0.76748 | 1.1588 | 0.66965 | 1.2566 | 10.16% | 20.31% | 30.47% | 0.91499 | 1.0113
In, ppm | 2.1391 | 0.089193 | 1.9607 | 2.3175 | 1.8715 | 2.4067 | 4.17% | 8.339% | 12.51% | 2.0321 | 2.246
K, wt.% | 2.5541 | 0.086366 | 2.3814 | 2.7269 | 2.295 | 2.8132 | 3.381% | 6.763% | 10.14% | 2.4264 | 2.6818
La, ppm | 68.2929 | 9.7774 | 48.7382 | 87.8476 | 38.9608 | 97.625 | 14.32% | 28.63% | 42.95% | 64.8783 | 71.7076
Li, ppm | 62.6357 | 2.3093 | 58.0172 | 67.2542 | 55.7079 | 69.5635 | 3.687% | 7.374% | 11.06% | 59.5039 | 65.7675
Lu, ppm | 0.41808 | 0.052038 | 0.31401 | 0.52216 | 0.26197 | 0.5742 | 12.45% | 24.89% | 37.34% | 0.39718 | 0.43899
Mg, wt.% | 3.5765 | 0.13132 | 3.3139 | 3.8391 | 3.1826 | 3.9705 | 3.672% | 7.343% | 11.01% | 3.3977 | 3.7553
Mn, wt.% | 0.013474 | 0.0002824 | 0.012909 | 0.014038 | 0.012626 | 0.014321 | 2.096% | 4.192% | 6.288% | 0.0128 | 0.014147
Mo, ppm | 15.0823 | 0.6663 | 13.7497 | 16.4149 | 13.0834 | 17.0812 | 4.418% | 8.836% | 13.25% | 14.3282 | 15.8364
Na, wt.% | 0.044918 | 0.0057524 | 0.033413 | 0.056423 | 0.02766 | 0.062175 | 12.81% | 25.61% | 38.42% | 0.042672 | 0.047164
Nb, ppm | 17.3826 | 3.0265 | 11.3297 | 23.4356 | 8.3032 | 26.4621 | 17.41% | 34.82% | 52.23% | 16.5135 | 18.2518
Nd, ppm | 62.7385 | 6.7681 | 49.2023 | 76.2747 | 42.4342 | 83.0428 | 10.79% | 21.58% | 32.36% | 59.6016 | 65.8754
Ni, ppm | 124.4568 | 3.4519 | 117.5529 | 131.3607 | 114.101 | 134.8126 | 2.774% | 5.547% | 8.321% | 118.234 | 130.6796
P, wt.% | 0.03505 | 0.0019227 | 0.031204 | 0.038895 | 0.029282 | 0.040818 | 5.486% | 10.97% | 16.46% | 0.033297 | 0.036802
Pb, ppm | 11.234 | 1.0643 | 9.1053 | 13.3627 | 8.041 | 14.427 | 9.474% | 18.95% | 28.42% | 10.6723 | 11.7957
Pr, ppm | 16.8542 | 1.2029 | 14.4483 | 19.2601 | 13.2453 | 20.463 | 7.137% | 14.27% | 21.41% | 16.0115 | 17.6969
Rb, ppm | 95.8242 | 5.0255 | 85.7732 | 105.8751 | 80.7477 | 110.9006 | 5.245% | 10.49% | 15.73% | 91.0329 | 100.6154
Re, ppm | 0.027479 | 0.0022045 | 0.02307 | 0.031888 | 0.020866 | 0.034093 | 8.022% | 16.04% | 24.07% | 0.026105 | 0.028853
S, wt.% | 3.3376 | 0.075088 | 3.1874 | 3.4878 | 3.1123 | 3.5629 | 2.25% | 4.5% | 6.749% | 3.1707 | 3.5045
Sb, ppm | 3.0122 | 0.34236 | 2.3275 | 3.6969 | 1.9851 | 4.0393 | 11.37% | 22.73% | 34.1% | 2.8616 | 3.1628
Sc, ppm | 12.5054 | 0.74634 | 11.0127 | 13.9981 | 10.2664 | 14.7444 | 5.968% | 11.94% | 17.9% | 11.8801 | 13.1307
Se, ppm | 5.729 | 0.52773 | 4.6735 | 6.7844 | 4.1458 | 7.3122 | 9.212% | 18.42% | 27.64% | 5.4425 | 6.0154
Sm, ppm | 8.6857 | 0.89449 | 6.8968 | 10.4747 | 6.0023 | 11.3692 | 10.3% | 20.6% | 30.9% | 8.2515 | 9.12
Sn, ppm | 2.8123 | 0.15853 | 2.4952 | 3.1293 | 2.3367 | 3.2879 | 5.637% | 11.27% | 16.91% | 2.6717 | 2.9529
Sr, ppm | 70.0375 | 2.8404 | 64.3567 | 75.7184 | 61.5163 | 78.5588 | 4.056% | 8.111% | 12.17% | 66.5357 | 73.5394
Ta, ppm | 1.2644 | 0.2046 | 0.85519 | 1.6736 | 0.65059 | 1.8782 | 16.18% | 32.36% | 48.55% | 1.2012 | 1.3276
Tb, ppm | 0.74168 | 0.088604 | 0.56447 | 0.91889 | 0.47587 | 1.0075 | 11.95% | 23.89% | 35.84% | 0.70459 | 0.77876
Te, ppm | 0.063233 | 0.010736 | 0.041761 | 0.084706 | 0.031025 | 0.095442 | 16.98% | 33.96% | 50.94% | 0.060072 | 0.066395
Th, ppm | 15.4502 | 0.61763 | 14.2149 | 16.6854 | 13.5973 | 17.3031 | 3.998% | 7.995% | 11.99% | 14.6777 | 16.2227
Ti, wt.% | 0.25811 | 0.070593 | 0.11692 | 0.3993 | 0.046332 | 0.46989 | 27.35% | 54.7% | 82.05% | 0.2452 | 0.27102
Tl, ppm | 0.66915 | 0.047639 | 0.57387 | 0.76443 | 0.52623 | 0.81207 | 7.119% | 14.24% | 21.36% | 0.63569 | 0.70261
Tm, ppm | 0.4214 | 0.075959 | 0.26948 | 0.57332 | 0.19352 | 0.64928 | 18.03% | 36.05% | 54.08% | 0.40033 | 0.44247
U, ppm | 8.6004 | 0.42071 | 7.759 | 9.4419 | 7.3383 | 9.8626 | 4.892% | 9.783% | 14.68% | 8.1704 | 9.0305
V, ppm | 348.4629 | 13.673 | 321.1169 | 375.8089 | 307.444 | 389.4819 | 3.924% | 7.848% | 11.77% | 331.0398 | 365.8861
W, ppm | 2.52 | 0.21493 | 2.0901 | 2.9498 | 1.8752 | 3.1648 | 8.529% | 17.06% | 25.59% | 2.394 | 2.646
Y, ppm | 23.7694 | 0.96611 | 21.8372 | 25.7016 | 20.8711 | 26.6677 | 4.064% | 8.129% | 12.19% | 22.5809 | 24.9579
Yb, ppm | 2.8396 | 0.40692 | 2.0257 | 3.6534 | 1.6188 | 4.0603 | 14.33% | 28.66% | 42.99% | 2.6976 | 2.9815
Zn, ppm | 16.1493 | 1.2088 | 13.7316 | 18.567 | 12.5228 | 19.7759 | 7.485% | 14.97% | 22.46% | 15.3419 | 16.9568
Zr, ppm | 187.8487 | 9.9615 | 167.9258 | 207.7717 | 157.9644 | 217.7331 | 5.303% | 10.61% | 15.91% | 178.4563 | 197.2412
Sulphuric Acid 5% Leach
Co, wt.% | 0.12698 | 0.0078932 | 0.11119 | 0.14276 | 0.1033 | 0.15066 | 6.216% | 12.43% | 18.65% | 0.12063 | 0.13332
Cu, wt.% | 1.0662 | 0.078623 | 0.90899 | 1.2235 | 0.83037 | 1.3021 | 7.374% | 14.75% | 22.12% | 1.0129 | 1.1195
Infrared Combustion
C, wt.% | 4.5091 | 0.13844 | 4.2322 | 4.786 | 4.0938 | 4.9244 | 3.07% | 6.14% | 9.21% | 4.2836 | 4.7345
S, wt.% | 3.402 | 0.0846 | 3.2328 | 3.5712 | 3.1482 | 3.6558 | 2.487% | 4.974% | 7.46% | 3.2319 | 3.5721
Borate Fusion XRF
Al2O3, wt.% | 12.5052 | 0.19688 | 12.1114 | 12.8989 | 11.9145 | 13.0958 | 1.574% | 3.149% | 4.723% | 11.8799 | 13.1304
BaO, ppm | 303.4666 | 44.6704 | 214.1258 | 392.8075 | 169.4553 | 437.478 | 14.72% | 29.44% | 44.16% | 288.2933 | 318.64
CaO, wt.% | 0.43499 | 0.014682 | 0.40563 | 0.46436 | 0.39095 | 0.47904 | 3.375% | 6.751% | 10.13% | 0.41324 | 0.45674
Co, wt.% | 0.58666 | 0.016772 | 0.55311 | 0.6202 | 0.53634 | 0.63697 | 2.859% | 5.718% | 8.577% | 0.55732 | 0.61599
Cu, wt.% | 3.6212 | 0.069282 | 3.4826 | 3.7597 | 3.4133 | 3.829 | 1.913% | 3.826% | 5.74% | 3.4401 | 3.8022
Fe2O3, wt.% | 4.5996 | 0.074463 | 4.4507 | 4.7485 | 4.3762 | 4.823 | 1.619% | 3.238% | 4.857% | 4.3696 | 4.8296
K2O, wt.% | 3.043 | 0.090923 | 2.8611 | 3.2248 | 2.7702 | 3.3158 | 2.988% | 5.976% | 8.964% | 2.8908 | 3.1951
MgO, wt.% | 5.9391 | 0.12665 | 5.6858 | 6.1924 | 5.5591 | 6.319 | 2.133% | 4.265% | 6.398% | 5.6421 | 6.236
MnO, wt.% | 0.016648 | 0.0039673 | 0.0087131 | 0.024582 | 0.0047457 | 0.02855 | 23.83% | 47.66% | 71.49% | 0.015815 | 0.01748
Ni, ppm | 125.2575 | 21.2053 | 82.847 | 167.668 | 61.6417 | 188.8732 | 16.93% | 33.86% | 50.79% | 118.9946 | 131.5204
P2O5, wt.% | 0.082926 | 0.0070297 | 0.068866 | 0.096985 | 0.061837 | 0.10401 | 8.477% | 16.95% | 25.43% | 0.078779 | 0.087072
SiO2, wt.% | 53.7231 | 0.58581 | 52.5514 | 54.8947 | 51.9656 | 55.4805 | 1.09% | 2.181% | 3.271% | 51.0369 | 56.4092
SO3, wt.% | 8.5942 | 0.21452 | 8.1651 | 9.0232 | 7.9506 | 9.2377 | 2.496% | 4.992% | 7.488% | 8.1645 | 9.0239
SrO, ppm | 95.6006 | 21.1345 | 53.3316 | 137.8696 | 32.1971 | 159.0041 | 22.11% | 44.21% | 66.32% | 90.8205 | 100.3806
TiO2, wt.% | 0.79148 | 0.015081 | 0.76132 | 0.82165 | 0.74624 | 0.83673 | 1.905% | 3.811% | 5.716% | 0.75191 | 0.83106
V2O5, ppm | 614.1301 | 76.8751 | 460.3799 | 767.8804 | 383.5047 | 844.7555 | 12.52% | 25.04% | 37.55% | 583.4236 | 644.8366
Zr, ppm | 240.4354 | 47.7817 | 144.8721 | 335.9988 | 97.0904 | 383.7804 | 19.87% | 39.75% | 59.62% | 228.4136 | 252.4572
Thermogravimetry
LOI1000, wt.% | 12.7221 | 0.34433 | 12.0334 | 13.4107 | 11.6891 | 13.755 | 2.707% | 5.413% | 8.12% | 12.086 | 13.3582
Peroxide Fusion ICP
Al, wt.% | 6.4666 | 0.15588 | 6.1548 | 6.7784 | 5.9989 | 6.9342 | 2.411% | 4.821% | 7.232% | 6.1433 | 6.7899
As, ppm | 52.866 | 6.3788 | 40.1085 | 65.6235 | 33.7297 | 72.0023 | 12.07% | 24.13% | 36.2% | 50.2227 | 55.5093
B, ppm | 217.1039 | 15.0739 | 186.9562 | 247.2516 | 171.8823 | 262.3255 | 6.943% | 13.89% | 20.83% | 206.2487 | 227.9591
Ba, ppm | 264.0065 | 10.0188 | 243.9689 | 284.044 | 233.9501 | 294.0628 | 3.795% | 7.59% | 11.38% | 250.8061 | 277.2068
Be, ppm | 5.4194 | 0.69589 | 4.0277 | 6.8112 | 3.3318 | 7.5071 | 12.84% | 25.68% | 38.52% | 5.1485 | 5.6904
Bi, ppm | 1.7722 | 0.11367 | 1.5449 | 1.9996 | 1.4312 | 2.1132 | 6.414% | 12.83% | 19.24% | 1.6836 | 1.8608
Ca, wt.% | 0.3067 | 0.037271 | 0.23216 | 0.38124 | 0.19488 | 0.41851 | 12.15% | 24.3% | 36.46% | 0.29136 | 0.32203
Ce, ppm | 151.6062 | 4.9578 | 141.6906 | 161.5218 | 136.7328 | 166.4797 | 3.27% | 6.54% | 9.811% | 144.0259 | 159.1865
Co, wt.% | 0.58122 | 0.023311 | 0.5346 | 0.62784 | 0.51129 | 0.65115 | 4.011% | 8.022% | 12.03% | 0.55216 | 0.61028
Cr, ppm | 110.467 | 16.3988 | 77.6694 | 143.2647 | 61.2705 | 159.6635 | 14.85% | 29.69% | 44.54% | 104.9437 | 115.9904
Cs, ppm | 2.1337 | 0.14093 | 1.8518 | 2.4156 | 1.7109 | 2.5565 | 6.605% | 13.21% | 19.81% | 2.027 | 2.2404
Cu, wt.% | 3.6061 | 0.11032 | 3.3854 | 3.8267 | 3.2751 | 3.9371 | 3.059% | 6.119% | 9.178% | 3.4258 | 3.7864
Dy, ppm | 7.0307 | 0.30361 | 6.4235 | 7.6379 | 6.1199 | 7.9415 | 4.318% | 8.637% | 12.96% | 6.6792 | 7.3822
Er, ppm | 3.7332 | 0.32945 | 3.0743 | 4.3921 | 2.7449 | 4.7216 | 8.825% | 17.65% | 26.47% | 3.5466 | 3.9199
Eu, ppm | 0.91026 | 0.094566 | 0.72113 | 1.0994 | 0.62656 | 1.194 | 10.39% | 20.78% | 31.17% | 0.86474 | 0.95577
Fe, wt.% | 3.1677 | 0.091679 | 2.9844 | 3.3511 | 2.8927 | 3.4428 | 2.894% | 5.788% | 8.682% | 3.0093 | 3.3261
Ga, ppm | 27.6572 | 1.0852 | 25.4868 | 29.8277 | 24.4015 | 30.9129 | 3.924% | 7.848% | 11.77% | 26.2744 | 29.0401
Gd, ppm | 6.8708 | 0.41871 | 6.0334 | 7.7083 | 5.6147 | 8.127 | 6.094% | 12.19% | 18.28% | 6.5273 | 7.2144
Ge, ppm | 2.2382 | 0.67391 | 0.89034 | 3.586 | 0.21643 | 4.2599 | 30.11% | 60.22% | 90.33% | 2.1263 | 2.3501
Ho, ppm | 1.352 | 0.10743 | 1.1371 | 1.5669 | 1.0297 | 1.6743 | 7.946% | 15.89% | 23.84% | 1.2844 | 1.4196
In, ppm | 2.2177 | 0.16646 | 1.8848 | 2.5506 | 1.7183 | 2.7171 | 7.506% | 15.01% | 22.52% | 2.1068 | 2.3286
K, wt.% | 2.5832 | 0.094236 | 2.3948 | 2.7717 | 2.3005 | 2.8659 | 3.648% | 7.296% | 10.94% | 2.4541 | 2.7124
La, ppm | 76.9894 | 4.1013 | 68.7868 | 85.192 | 64.6855 | 89.2933 | 5.327% | 10.65% | 15.98% | 73.14 | 80.8389
Li, ppm | 64.8582 | 4.5767 | 55.7048 | 74.0116 | 51.1281 | 78.5883 | 7.056% | 14.11% | 21.17% | 61.6153 | 68.1011
Lu, ppm | 0.55961 | 0.050431 | 0.45874 | 0.66047 | 0.40831 | 0.7109 | 9.012% | 18.02% | 27.04% | 0.53163 | 0.58759
Mg, wt.% | 3.5675 | 0.086569 | 3.3944 | 3.7406 | 3.3078 | 3.8272 | 2.427% | 4.853% | 7.28% | 3.3891 | 3.7459
Mn, wt.% | 0.014047 | 0.0016582 | 0.010731 | 0.017364 | 0.0090728 | 0.019022 | 11.8% | 23.61% | 35.41% | 0.013345 | 0.01475
Mo, ppm | 15.9784 | 1.4737 | 13.031 | 18.9258 | 11.5573 | 20.3995 | 9.223% | 18.45% | 27.67% | 15.1795 | 16.7773
Nb, ppm | 25.5676 | 1.5838 | 22.4001 | 28.7351 | 20.8163 | 30.3188 | 6.194% | 12.39% | 18.58% | 24.2892 | 26.846
Nd, ppm | 64.6418 | 2.0947 | 60.4524 | 68.8313 | 58.3576 | 70.926 | 3.241% | 6.481% | 9.722% | 61.4097 | 67.8739
Ni, ppm | 125.636 | 11.4657 | 102.7045 | 148.5675 | 91.2388 | 160.0332 | 9.126% | 18.25% | 27.38% | 119.3542 | 131.9178
P, wt.% | 0.037287 | 0.0076346 | 0.022017 | 0.052556 | 0.014383 | 0.06019 | 20.48% | 40.95% | 61.43% | 0.035422 | 0.039151
Pb, ppm | 32.8276 | 19.5408 | 0 | 71.9091 | 0 | 91.4499 | 59.53% | 119.1% | 178.6% | 31.1862 | 34.4689
Pr, ppm | 18.131 | 0.5055 | 17.12 | 19.142 | 16.6145 | 19.6475 | 2.788% | 5.576% | 8.364% | 17.2244 | 19.0375
Rb, ppm | 96.2434 | 4.0608 | 88.1217 | 104.3651 | 84.0608 | 108.4259 | 4.219% | 8.439% | 12.66% | 91.4312 | 101.0555
Re, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 3.2919 | 0.18772 | 2.9164 | 3.6673 | 2.7287 | 3.8551 | 5.703% | 11.41% | 17.11% | 3.1273 | 3.4565
Sb, ppm | 3.3495 | 0.27365 | 2.8022 | 3.8968 | 2.5286 | 4.1705 | 8.17% | 16.34% | 24.51% | 3.1821 | 3.517
Sc, ppm | 12.2481 | 0.99226 | 10.2636 | 14.2327 | 9.2714 | 15.2249 | 8.101% | 16.2% | 24.3% | 11.6357 | 12.8606
Si, wt.% | 25.1713 | 0.78896 | 23.5934 | 26.7493 | 22.8044 | 27.5382 | 3.134% | 6.269% | 9.403% | 23.9128 | 26.4299
Sm, ppm | 8.9016 | 0.66278 | 7.576 | 10.2271 | 6.9132 | 10.8899 | 7.446% | 14.89% | 22.34% | 8.4565 | 9.3466
Sr, ppm | 69.6557 | 3.6363 | 62.383 | 76.9284 | 58.7467 | 80.5648 | 5.22% | 10.44% | 15.66% | 66.173 | 73.1385
Ta, ppm | 2.0098 | 0.29629 | 1.4173 | 2.6024 | 1.121 | 2.8987 | 14.74% | 29.48% | 44.23% | 1.9093 | 2.1103
Tb, ppm | 1.1832 | 0.067539 | 1.0481 | 1.3182 | 0.98054 | 1.3858 | 5.708% | 11.42% | 17.13% | 1.124 | 1.2423
Th, ppm | 15.9533 | 0.39287 | 15.1676 | 16.7391 | 14.7747 | 17.1319 | 2.463% | 4.925% | 7.388% | 15.1557 | 16.751
Ti, wt.% | 0.46085 | 0.017072 | 0.42671 | 0.495 | 0.40964 | 0.51207 | 3.704% | 7.409% | 11.11% | 0.43781 | 0.4839
Tl, ppm | 0.7292 | 0.064881 | 0.59944 | 0.85896 | 0.53456 | 0.92385 | 8.897% | 17.79% | 26.69% | 0.69274 | 0.76566
Tm, ppm | 0.57021 | 0.051756 | 0.4667 | 0.67372 | 0.41494 | 0.72548 | 9.077% | 18.15% | 27.23% | 0.5417 | 0.59872
U, ppm | 8.6995 | 0.41125 | 7.877 | 9.522 | 7.4658 | 9.9333 | 4.727% | 9.455% | 14.18% | 8.2645 | 9.1345
V, ppm | 364.947 | 15.9545 | 333.0381 | 396.8559 | 317.0836 | 412.8104 | 4.372% | 8.743% | 13.12% | 346.6996 | 383.1943
W, ppm | 3.1755 | 0.49492 | 2.1857 | 4.1653 | 1.6908 | 4.6603 | 15.59% | 31.17% | 46.76% | 3.0167 | 3.3343
Y, ppm | 35.6106 | 1.892 | 31.8267 | 39.3946 | 29.9347 | 41.2866 | 5.313% | 10.63% | 15.94% | 33.8301 | 37.3912
Yb, ppm | 3.6337 | 0.34826 | 2.9372 | 4.3302 | 2.589 | 4.6785 | 9.584% | 19.17% | 28.75% | 3.452 | 3.8154
Aqua Regia Digestion
Ag, ppm | 0.076744 | 0.0084537 | 0.059836 | 0.093651 | 0.051382 | 0.1021 | 11.02% | 22.03% | 33.05% | 0.072906 | 0.080581
Al, wt.% | 0.93966 | 0.13478 | 0.67011 | 1.2092 | 0.53533 | 1.344 | 14.34% | 28.69% | 43.03% | 0.89268 | 0.98664
As, ppm | 51.3639 | 2.5265 | 46.3109 | 56.4169 | 43.7844 | 58.9434 | 4.919% | 9.838% | 14.76% | 48.7957 | 53.9321
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 49.8917 | 6.0618 | 37.7682 | 62.0152 | 31.7064 | 68.077 | 12.15% | 24.3% | 36.45% | 47.3971 | 52.3863
Be, ppm | 2.1563 | 0.18888 | 1.7785 | 2.534 | 1.5896 | 2.7229 | 8.76% | 17.52% | 26.28% | 2.0485 | 2.2641
Bi, ppm | 1.5791 | 0.1108 | 1.3575 | 1.8007 | 1.2467 | 1.9115 | 7.017% | 14.03% | 21.05% | 1.5002 | 1.6581
Ca, wt.% | 0.30016 | 0.013306 | 0.27355 | 0.32677 | 0.26024 | 0.34008 | 4.433% | 8.866% | 13.3% | 0.28515 | 0.31517
Cd, ppm | 0.29845 | 0.023467 | 0.25151 | 0.34538 | 0.22805 | 0.36885 | 7.863% | 15.73% | 23.59% | 0.28353 | 0.31337
Ce, ppm | 42.8634 | 6.2347 | 30.394 | 55.3327 | 24.1593 | 61.5674 | 14.55% | 29.09% | 43.64% | 40.7202 | 45.0066
Co, wt.% | 0.54777 | 0.053332 | 0.44111 | 0.65443 | 0.38777 | 0.70776 | 9.736% | 19.47% | 29.21% | 0.52038 | 0.57516
Cr, ppm | 33.9115 | 3.313 | 27.2855 | 40.5375 | 23.9725 | 43.8505 | 9.77% | 19.54% | 29.31% | 32.2159 | 35.6071
Cs, ppm | 0.29441 | 0.079683 | 0.13504 | 0.45378 | 0.055359 | 0.53346 | 27.07% | 54.13% | 81.2% | 0.27969 | 0.30913
Cu, wt.% | 3.6039 | 0.077422 | 3.4491 | 3.7587 | 3.3716 | 3.8362 | 2.148% | 4.297% | 6.445% | 3.4237 | 3.7841
Dy, ppm | 0.79906 | 0.083738 | 0.63158 | 0.96654 | 0.54785 | 1.0503 | 10.48% | 20.96% | 31.44% | 0.75911 | 0.83901
Er, ppm | 0.36324 | 0.029655 | 0.30393 | 0.42255 | 0.27427 | 0.4522 | 8.164% | 16.33% | 24.49% | 0.34508 | 0.3814
Eu, ppm | 0.31443 | 0.035345 | 0.24374 | 0.38512 | 0.20839 | 0.42046 | 11.24% | 22.48% | 33.72% | 0.2987 | 0.33015
Fe, wt.% | 2.9675 | 0.11125 | 2.745 | 3.1901 | 2.6338 | 3.3013 | 3.749% | 7.498% | 11.25% | 2.8192 | 3.1159
Ga, ppm | 7.6749 | 1.294 | 5.087 | 10.2629 | 3.7931 | 11.5568 | 16.86% | 33.72% | 50.58% | 7.2912 | 8.0587
Gd, ppm | 1.7358 | 0.22604 | 1.2837 | 2.1879 | 1.0577 | 2.4139 | 13.02% | 26.04% | 39.07% | 1.649 | 1.8226
Hf, ppm | 0.30042 | 0.023505 | 0.25341 | 0.34743 | 0.22991 | 0.37093 | 7.824% | 15.65% | 23.47% | 0.2854 | 0.31544
Hg, ppm | 0.050371 | 0.012332 | 0.025707 | 0.075035 | 0.013376 | 0.087367 | 24.48% | 48.96% | 73.45% | 0.047853 | 0.05289
Ho, ppm | 0.1392 | 0.026268 | 0.086662 | 0.19174 | 0.060394 | 0.218 | 18.87% | 37.74% | 56.61% | 0.13224 | 0.14616
In, ppm | 1.5437 | 0.10837 | 1.327 | 1.7605 | 1.2186 | 1.8688 | 7.02% | 14.04% | 21.06% | 1.4665 | 1.6209
K, wt.% | 0.19011 | 0.016757 | 0.15659 | 0.22362 | 0.13983 | 0.24038 | 8.815% | 17.63% | 26.44% | 0.1806 | 0.19961
La, ppm | 18.9644 | 2.4178 | 14.1287 | 23.8001 | 11.7108 | 26.2179 | 12.75% | 25.5% | 38.25% | 18.0161 | 19.9126
Li, ppm | 35.4644 | 6.2661 | 22.9322 | 47.9966 | 16.666 | 54.2627 | 17.67% | 35.34% | 53.01% | 33.6912 | 37.2376
Lu, ppm | 0.041033 | 0.002498 | 0.036037 | 0.046029 | 0.033539 | 0.048527 | 6.088% | 12.18% | 18.26% | 0.038982 | 0.043085
Mg, wt.% | 2.8486 | 0.2086 | 2.4314 | 3.2658 | 2.2228 | 3.4744 | 7.323% | 14.65% | 21.97% | 2.7062 | 2.991
Mn, wt.% | 0.012506 | 0.00076284 | 0.01098 | 0.014032 | 0.010217 | 0.014795 | 6.1% | 12.2% | 18.3% | 0.011881 | 0.013131
Mo, ppm | 14.1916 | 0.60889 | 12.9739 | 15.4094 | 12.365 | 16.0183 | 4.29% | 8.581% | 12.87% | 13.4821 | 14.9012
Na, wt.% | 0.0097809 | 0.00073018 | 0.0083205 | 0.011241 | 0.0075903 | 0.011971 | 7.465% | 14.93% | 22.4% | 0.0092918 | 0.01027
Nd, ppm | 20.1566 | 3.7281 | 12.7005 | 27.6128 | 8.9724 | 31.3408 | 18.5% | 36.99% | 55.49% | 19.1488 | 21.1644
Ni, ppm | 115.4174 | 4.4086 | 106.6001 | 124.2346 | 102.1915 | 128.6432 | 3.82% | 7.639% | 11.46% | 109.6465 | 121.1882
P, wt.% | 0.020547 | 0.001386 | 0.017775 | 0.023319 | 0.016389 | 0.024705 | 6.745% | 13.49% | 20.24% | 0.019519 | 0.021574
Pb, ppm | 8.42 | 0.83735 | 6.7453 | 10.0947 | 5.908 | 10.9321 | 9.945% | 19.89% | 29.83% | 7.999 | 8.841
Pr, ppm | 5.2784 | 0.89104 | 3.4963 | 7.0605 | 2.6053 | 7.9516 | 16.88% | 33.76% | 50.64% | 5.0145 | 5.5423
Rb, ppm | 6.2944 | 1.1403 | 4.0138 | 8.5749 | 2.8735 | 9.7152 | 18.12% | 36.23% | 54.35% | 5.9796 | 6.6091
Re, ppm | 0.025986 | 0.0014642 | 0.023058 | 0.028915 | 0.021594 | 0.030379 | 5.634% | 11.27% | 16.9% | 0.024687 | 0.027286
S, wt.% | 3.2631 | 0.10348 | 3.0561 | 3.4701 | 2.9526 | 3.5735 | 3.171% | 6.343% | 9.514% | 3.0999 | 3.4262
Sb, ppm | 1.8554 | 0.22769 | 1.4 | 2.3108 | 1.1723 | 2.5385 | 12.27% | 24.54% | 36.82% | 1.7626 | 1.9482
Sc, ppm | 3.2999 | 0.35987 | 2.5801 | 4.0196 | 2.2203 | 4.3795 | 10.91% | 21.81% | 32.72% | 3.1349 | 3.4649
Se, ppm | 5.6527 | 0.36688 | 4.919 | 6.3865 | 4.5521 | 6.7534 | 6.49% | 12.98% | 19.47% | 5.3701 | 5.9354
Sm, ppm | 3.197 | 0.46258 | 2.2718 | 4.1221 | 1.8092 | 4.5847 | 14.47% | 28.94% | 43.41% | 3.0371 | 3.3568
Sn, ppm | 1.2775 | 0.069802 | 1.1379 | 1.4171 | 1.0681 | 1.4869 | 5.464% | 10.93% | 16.39% | 1.2136 | 1.3413
Sr, ppm | 16.4063 | 2.4687 | 11.4689 | 21.3436 | 9.0002 | 23.8123 | 15.05% | 30.09% | 45.14% | 15.5859 | 17.2266
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.19025 | 0.014943 | 0.16037 | 0.22014 | 0.14542 | 0.23508 | 7.854% | 15.71% | 23.56% | 0.18074 | 0.19977
Te, ppm | 0.047667 | 0.01164 | 0.024386 | 0.070948 | 0.012746 | 0.082588 | 24.42% | 48.84% | 73.26% | 0.045284 | 0.05005
Th, ppm | 6.1773 | 0.66472 | 4.8479 | 7.5067 | 4.1832 | 8.1715 | 10.76% | 21.52% | 32.28% | 5.8684 | 6.4862
U, ppm | 3.0052 | 0.22559 | 2.554 | 3.4563 | 2.3284 | 3.6819 | 7.507% | 15.01% | 22.52% | 2.8549 | 3.1554
V, ppm | 55.7682 | 8.5776 | 38.613 | 72.9234 | 30.0354 | 81.501 | 15.38% | 30.76% | 46.14% | 52.9798 | 58.5566
W, ppm | 0.96473 | 0.066843 | 0.83104 | 1.0984 | 0.7642 | 1.1653 | 6.929% | 13.86% | 20.79% | 0.91649 | 1.013
Y, ppm | 3.0963 | 0.37588 | 2.3446 | 3.8481 | 1.9687 | 4.224 | 12.14% | 24.28% | 36.42% | 2.9415 | 3.2512
Yb, ppm | 0.28194 | 0.038594 | 0.20475 | 0.35912 | 0.16615 | 0.39772 | 13.69% | 27.38% | 41.07% | 0.26784 | 0.29603
Zn, ppm | 13.0011 | 1.2315 | 10.5381 | 15.4641 | 9.3066 | 16.6956 | 9.472% | 18.94% | 28.42% | 12.3511 | 13.6512
Zr, ppm | 9.3918 | 1.3526 | 6.6866 | 12.097 | 5.334 | 13.4496 | 14.4% | 28.8% | 43.21% | 8.9222 | 9.8614
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 553b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AL-5S*AAS | 5% sulphuric acid laech with atomic absorption spectroscopy
AL-5S*OES | 5% sulphuric acid leach with inductively coupled plasma optical emission spectroscopy
AL-5S*OES/AAS | 5% sulphuric acid leach with AAS or ICP-OES finish
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Text Values:
NA | Unable to report due to QC failure (Lab 19)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 553b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Labwest Minerals Analysis, Perth, WA, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 553b (Certified Value 0.103 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.1 | 0.196 | 0.5 | < 0.1 | 1 | 0.11 | < 0.1 | 0.09 | 0.11 | 0.09 | 0.09 | 0.08 | 0.14505 | 0.2 | 0.11 | 0.15 | < 0.3 | 0.1 | 0.1 | 0.1 | 0.12 |  |  | 1
| 1 | 2 | 0.09 | 0.11 | 0.4 | 0.1 | 1 | 0.09 | < 0.1 | 0.1 | 0.07 | 0.11 | 0.11 | 0.08 | 0.14492 | 0.2 | 0.11 | 0.14 | < 0.3 | 0.1 | 0.1 | 0.11 | 0.1 |  |  | 4
| 1 | 3 | 0.09 | 0.123 | 0.5 | < 0.1 | 1 | 0.1 | < 0.1 | 0.08 | 0.12 | 0.11 | 0.11 | 0.09 | 0.15059 | 0.3 | 0.11 | 0.16 | < 0.3 | 0.1 | 0.1 | 0.11 | 0.08 |  |  | 16
| 1 | 4 | 0.1 | 0.101 | 0.4 | 0.1 | 1 | 0.09 | < 0.1 | 0.11 | 0.1 | 0.09 | 0.09 | 0.08 | 0.15385 | 0.2 | 0.12 | 0.13 | < 0.3 | 0.1 | < 0.1 | 0.12 | 0.1 |  |  | 0.10281
| 1 | 5 | 0.09 | 0.107 | 0.5 | 0.1 | 1 | 0.11 | < 0.1 | 0.1 | 0.12 | 0.1 | 0.09 | 0.1 | 0.14648 | 0.3 | 0.11 | 0.14 | < 0.3 | 0.1 | < 0.1 | 0.13 | 0.11 |  |  | 7
| 1 | 6 | 0.1 | 0.12 | 0.5 | 0.1 | 1 | 0.1 | < 0.1 | 0.13 | 0.11 | 0.1 | 0.1 | 0.09 | 0.14242 | 0.2 | 0.11 | 0.15 | < 0.3 | 0.1 | < 0.1 | 0.1 | 0.1
| Mean |  | 0.095 | 0.12617 | 0.46667 | 0.1 | 1 | 0.1 |  | 0.10167 | 0.105 | 0.1 | 0.098333 | 0.086667 | 0.14722 | 0.23333 | 0.11167 | 0.145 |  | 0.1 | 0.1 | 0.11167 | 0.10167
| Median |  | 0.095 | 0.115 | 0.5 | 0.1 | 1 | 0.1 |  | 0.1 | 0.11 | 0.1 | 0.095 | 0.085 | 0.14576 | 0.2 | 0.11 | 0.145 |  | 0.1 | 0.1 | 0.11 | 0.1
| Std Dev. |  | 0.0054772 | 0.035176 | 0.05164 | 0 | 0 | 0.0089443 |  | 0.017224 | 0.018708 | 0.0089443 | 0.0098319 | 0.008165 | 0.0042175 | 0.05164 | 0.0040825 | 0.010488 |  | 1.5202e-17 | 1.6997e-17 | 0.01169 | 0.013292
| Rel.Std.Dev. |  | 5.766% | 27.88% | 11.07% | 0% | 0% | 8.944% |  | 16.94% | 17.82% | 8.944% | 9.999% | 9.421% | 2.865% | 22.13% | 3.656% | 7.233% |  | 1.52e-14% | 1.7e-14% | 10.47% | 13.07%
| PDM3 |  | -7.593% | 22.72% | 353.9% | -2.729% | 872.7% | -2.729% |  | -1.108% | 2.134% | -2.729% | -4.351% | -15.7% | 43.2% | 127% | 8.619% | 41.04% |  | -2.729% | -2.729% | 8.619% | -1.108%
| Z-Score (Absolute) |  | 0.45 | 1.65 | NA | NA | NA | 0.11 | 3.48 | 0 | 0.22 | 0.11 | 0.22 | 1.01 | 3.07 | NA | 0.67 | 2.92 | 3.26 | NA | NA | 0.67 | 0
Results from laboratories 3, 4, 7, 16, 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 553b (Certified Value 6.41 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 6.4606 | 6.18 | 6.24 | 6.64 | 6.4388 | 7.31 | 6.54 | 6.26 | 6.4531 | 6.49 | 6.47 | 6.37 | 6.61 | 6.6 | 6.5313 | 6.44 | 6.18 | 6.7363 | 6.4835 | 5.99 | 6.42 | 6.18 | 5.65 | 1
| 1 | 2 | 6.4246 | 6.22 | 6.2 | 6.45 | 6.4338 | 6.98 | 6.52 | 6.16 | 6.4583 | 6.47 | 6.51 | 6.27 | 6.55 | 6.47 | 6.6422 | 6.37 | 6.23 | 6.7844 | 6.3786 | 6.05 | 6.49 | 6.32 | 5.77 | #N/A
| 1 | 3 | 6.4847 | 6.3 | 6.12 | 6.36 | 6.4445 | 7.14 | 6.56 | 6.31 | 6.4569 | 6.55 | 6.27 | 6.3 | 6.66 | 6.5 | 6.6502 | 6.42 | 6.18 | 6.7771 | 6.4307 | 6.05 | 6.41 | 6.41 | 5.54 | 16
| 1 | 4 | 6.598 | 6.07 | 6.26 | 6.36 | 6.4317 | 6.49 | 6.54 | 6.13 | 6.483 | 6.44 | 6.12 | 6.37 | 6.77 | 6.48 | 6.7066 | 6.44 | 6.25 | 6.8388 | 6.1894 | 6.13 | 6.46 | 6.21 | 5.61 | 6.4086
| 1 | 5 | 6.6669 | 6.92 | 6.18 | 6.62 | 6.4441 | 5.81 | 6.45 | 6.32 | 6.4421 | 6.46 | 6.33 | 6.3 | 6.73 | 6.52 | 6.6269 | 6.46 | 6.27 | 6.8127 | 6.2487 | 6.12 | 6.49 | 6 | 5.1 | 8
| 1 | 6 | 6.48 | 6.35 | 6.23 | 6.42 | 6.4313 | 5.92 | 6.45 | 6.2 | 6.4841 | 6.64 | 6.49 | 6.28 | 6.93 | 6.47 | 6.4941 | 6.45 | 6.19 | 6.8937 | 6.4722 | 6.39 | 6.39 | 6.25 | 4.71
| Mean |  | 6.5191 | 6.34 | 6.205 | 6.475 | 6.4374 | 6.6083 | 6.51 | 6.23 | 6.4629 | 6.5083 | 6.365 | 6.315 | 6.7083 | 6.5067 | 6.6085 | 6.43 | 6.2167 | 6.8072 | 6.3672 | 6.1217 | 6.4433 | 6.2283 | 5.3967
| Median |  | 6.4824 | 6.26 | 6.215 | 6.435 | 6.4363 | 6.735 | 6.53 | 6.23 | 6.4576 | 6.48 | 6.4 | 6.3 | 6.695 | 6.49 | 6.6346 | 6.44 | 6.21 | 6.7985 | 6.4047 | 6.085 | 6.44 | 6.23 | 5.575
| Std Dev. |  | 0.092881 | 0.30033 | 0.050498 | 0.12518 | 0.0060123 | 0.63848 | 0.048166 | 0.078994 | 0.016977 | 0.074677 | 0.15359 | 0.044159 | 0.13452 | 0.049666 | 0.07985 | 0.032249 | 0.038816 | 0.054706 | 0.12196 | 0.1412 | 0.04274 | 0.13906 | 0.40722
| Rel.Std.Dev. |  | 1.425% | 4.737% | 0.8138% | 1.933% | 0.0934% | 9.662% | 0.7399% | 1.268% | 0.2627% | 1.147% | 2.413% | 0.6993% | 2.005% | 0.7633% | 1.208% | 0.5015% | 0.6244% | 0.8037% | 1.915% | 2.307% | 0.6633% | 2.233% | 7.546%
| PDM3 |  | 1.724% | -1.071% | -3.177% | 1.036% | 0.4485% | 3.116% | 1.582% | -2.787% | 0.8471% | 1.556% | -0.6806% | -1.461% | 4.677% | 1.53% | 3.12% | 0.3336% | -2.995% | 6.219% | -0.6466% | -4.478% | 0.5417% | -2.813% | -15.79%
| Z-Score (Absolute) |  | 0.57 | 0.67 | 1.61 | 0.26 | 0 | 1.18 | 0.5 | 1.44 | 0.18 | 0.49 | 0.5 | 0.85 | 1.88 | 0.48 | 1.19 | 0.05 | 1.53 | 2.56 | 0.49 | 2.19 | 0.04 | 1.45 | 7.21
| Analytical results for As in OREAS 553b (Certified Value 54 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
1
| 1 | 1 | 51.4 | 47.6 | 56.2 | 53 | 53 | 58 | 51.1948 | 52.5 | 58.1 | 55.9 | 52.9 | 55.7 | 56.4559 | 52.83 | 54 | 53.7 | 57.1 | 28 | 30.5 | 50.9 | 56 |  |  | 1
| 1 | 2 | 54.3 | 47.9 | 57 | 54 | 54 | 54 | 49.7012 | 51.2 | 61.6 | 54.8 | 53.1 | 52.1 | 57.5172 | 51.88 | 54 | 56.3 | 55.8 | 29 | 32.9 | 48.6 | 54 |  |  | 5
| 1 | 3 | 52.7 | 48 | 56.9 | 52 | 53 | 56 | 51.0486 | 51.6 | 55.8 | 54.5 | 51.1 | 52.4 | 57.6069 | 51.67 | 54 | 54 | 55.5 | 30 | 31.5 | 49.3 | 49 |  |  | 16
| 1 | 4 | 50.8 | 48 | 60.4 | 54 | 54 | 55 | 49.1764 | 52.7 | 57.2 | 54.5 | 53.5 | 53 | 56.4258 | 52.85 | 55 | 54.4 | 51.9 | 31 | 32.2 | 51.6 | 52 |  |  | 53.5899
| 1 | 5 | 53.2 | 51.2 | 58.3 | 53 | 53 | 57 | 48.5177 | 50.5 | 57.8 | 55.7 | 54.1 | 52.4 | 56.4908 | 52.25 | 54 | 55.9 | 54.7 | 29 | 31.8 | 49 | 56 |  |  | 9
| 1 | 6 | 52.1 | 47.7 | 57.4 | 54 | 54 | 54 | 50.7073 | 53.6 | 59.7 | 55 | 55.6 | 49.8 | 56.1472 | 51.78 | 54 | 56.5 | 56.6 | 30 | 32.4 | 46.7 | 50
| Mean |  | 52.4167 | 48.4 | 57.7 | 53.3333 | 53.5 | 55.6667 | 50.0577 | 52.0167 | 58.3667 | 55.0667 | 53.3833 | 52.5667 | 56.774 | 52.21 | 54.1667 | 55.1333 | 55.2667 | 29.5 | 31.8833 | 49.35 | 52.8333
| Median |  | 52.4 | 47.95 | 57.2 | 53.5 | 53.5 | 55.5 | 50.2043 | 52.05 | 57.95 | 54.9 | 53.3 | 52.4 | 56.4733 | 52.065 | 54 | 55.15 | 55.65 | 29.5 | 32 | 49.15 | 53
| Std Dev. |  | 1.264 | 1.3813 | 1.4913 | 0.8165 | 0.54772 | 1.633 | 1.0929 | 1.1268 | 2.0285 | 0.60222 | 1.4811 | 1.8938 | 0.62316 | 0.52555 | 0.40825 | 1.2404 | 1.8511 | 1.0488 | 0.83287 | 1.7421 | 2.9944
| Rel.Std.Dev. |  | 2.411% | 2.854% | 2.585% | 1.531% | 1.024% | 2.934% | 2.183% | 2.166% | 3.475% | 1.094% | 2.774% | 3.603% | 1.098% | 1.007% | 0.7537% | 2.25% | 3.349% | 3.555% | 2.612% | 3.53% | 5.668%
| PDM3 |  | -2.189% | -9.684% | 7.67% | -0.4788% | -0.1678% | 3.875% | -6.591% | -2.936% | 8.914% | 2.756% | -0.3855% | -1.909% | 5.942% | -2.575% | 1.076% | 2.88% | 3.129% | -44.95% | -40.5% | -7.912% | -1.412%
| Z-Score (Absolute) |  | 0.34 | 1.85 | 1.64 | 0 | 0.06 | 0.87 | 1.23 | 0.49 | 1.89 | 0.65 | 0.02 | 0.29 | 1.29 | 0.42 | 0.31 | 0.67 | 0.72 | 8.93 | 8.04 | 1.49 | 0.19
| Analytical results for B in OREAS 553b (Indicative Value 20 ppm) | [24 empty] | Indicative
B | Round | Replicate | Lab | [22 empty] | 1
| No | No | 07 | [22 empty] | ppm
|  |  | 4A*OES | [22 empty] | 1
1
| 1 | 1 | 20 | [22 empty] | 1
| 1 | 2 | 20 | [22 empty] | 1
| 1 | 3 | 20 | [22 empty] | 16
| 1 | 4 | 20 | [22 empty] | 20
| 1 | 5 | 20 | [22 empty] | 7
| 1 | 6 | 20
| Mean |  | 20
| Median |  | 20
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 553b (Certified Value 262 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
0
| 1 | 1 | 275.4 | 71 | 270 | 261 | 342 | 264 | 243 | 234.2422 | 270 | 270 | 260 | 280 | 270 | 256.7039 | 255 | 285.5 | 257 | 153 | 274 | 190 | 240 |  |  | 1
| 1 | 2 | 274.7 | 39 | 268 | 264 | 317 | 263 | 256 | 233.8983 | 260 | 270 | 260 | 280 | 270 | 253.0216 | 255 | 282.3 | 253 | 175 | 275 | 194 | 238 |  |  | #N/A
| 1 | 3 | 275.7 | 55 | 275 | 264 | 336 | 265 | 254 | 234.7549 | 270 | 260 | 260 | 280 | 270 | 260.8156 | 255 | 283 | 254 | 181 | 272 | 187 | 229 |  |  | 16
| 1 | 4 | 275.7 | 82 | 269 | 260 | 299 | 262 | 241 | 240.7555 | 260 | 250 | 260 | 280 | 270 | 257.2936 | 258 | 280 | 248 | 206 | 273 | 175 | 231 |  |  | 261.8487
| 1 | 5 | 278.8 | 42 | 248 | 263 | 271 | 263 | 253 | 238.0359 | 260 | 260 | 260 | 280 | 270 | 253.5256 | 257 | 281.5 | 249 | 169 | 272 | 184 | 210 |  |  | 10
| 1 | 6 | 280.4 | 42 | 274 | 266 | 270 | 263 | 242 | 241.0908 | 270 | 270 | 260 | 290 | 270 | 252.8266 | 256 | 285.6 | 258 | 160 | 274 | 184 | 173
| Mean |  | 276.7833 | 55.1667 | 267.3333 | 263 | 305.8333 | 263.3333 | 248.1667 | 237.1296 | 265 | 263.3333 | 260 | 281.6667 | 270 | 255.6978 | 256 | 282.9833 | 253.1667 | 174 | 273.3333 | 185.6667 | 220.1667
| Median |  | 275.7 | 48.5 | 269.5 | 263.5 | 308 | 263 | 248 | 236.3954 | 265 | 265 | 260 | 280 | 270 | 255.1147 | 255.5 | 282.65 | 253.5 | 172 | 273.5 | 185.5 | 230
| Std Dev. |  | 2.2693 | 17.7698 | 9.8725 | 2.1909 | 31.2629 | 1.0328 | 6.8532 | 3.2887 | 5.4772 | 8.165 | 0 | 4.0825 | 0 | 3.1583 | 1.2649 | 2.2248 | 4.0702 | 18.6333 | 1.2111 | 6.4704 | 25.4355
| Rel.Std.Dev. |  | 0.8199% | 32.21% | 3.693% | 0.833% | 10.22% | 0.3922% | 2.762% | 1.387% | 2.067% | 3.101% | 0% | 1.449% | 0% | 1.235% | 0.4941% | 0.7862% | 1.608% | 10.71% | 0.4431% | 3.485% | 11.55%
| PDM3 |  | 5.704% | -78.93% | 2.095% | 0.4397% | 16.8% | 0.567% | -5.225% | -9.44% | 1.203% | 0.567% | -0.706% | 7.568% | 3.113% | -2.349% | -2.234% | 8.071% | -3.316% | -33.55% | 4.386% | -29.09% | -15.92%
| Z-Score (Absolute) |  | 0.9 | 13.56 | 0.28 | 0 | 2.8 | 0.02 | 0.97 | 1.69 | 0.13 | 0.02 | 0.2 | 1.22 | 0.46 | 0.48 | 0.46 | 1.3 | 0.64 | 5.81 | 0.67 | 5.05 | 2.8
| Analytical results for Be in OREAS 553b (Certified Value 4.71 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.82 | 6 | 5.5 | 4.6 | 10.567 | 4.94 | 3.8 | 4.8 | 4.9752 | 4.62 | 4.7 | 4.68 | 4.61 | 4.74 | 4.5473 | 4.95 | 4.48 | 4.83 | 3.99 | 5 | 5 | 4.9 | 5.5 | 1
| 1 | 2 | 4.85 | 5 | 5.6 | 4.5 | 10.5575 | 4.4 | 3.5 | 4.5 | 5.1803 | 4.55 | 4.7 | 4.66 | 4.55 | 4.64 | 4.6782 | 4.68 | 4.54 | 4.85 | 3.97 | 5 | 5 | 4.87 | 5.5 | 6
| 1 | 3 | 4.98 | 5 | 5.4 | 4.5 | 10.594 | 4.67 | 3.7 | 4.7 | 5.0379 | 4.57 | 4.55 | 4.59 | 4.49 | 4.65 | 4.7223 | 4.95 | 4.69 | 4.91 | 3.98 | 4 | 5 | 4.34 | 5.5 | 16
| 1 | 4 | 5.08 | 5 | 5.9 | 4.5 | 10.371 | 4.39 | 3.6 | 4.5 | 4.7218 | 4.54 | 4.62 | 4.66 | 4.58 | 4.65 | 4.8806 | 4.82 | 4.76 | 4.81 | 3.85 | 4 | 5 | 5.2 | 5.6 | 4.7054
| 1 | 5 | 5.2 | 6 | 5.5 | 4.5 | 10.596 | 4.53 | 3.6 | 4.8 | 4.7089 | 4.6 | 4.59 | 4.6 | 4.66 | 4.67 | 4.8017 | 4.82 | 4.67 | 4.85 | 3.89 | 4 | 5 | 4.23 | 5.5 | 11
| 1 | 6 | 4.97 | 5 | 5.5 | 4.6 | 10.326 | 4.38 | 3.7 | 4.6 | 4.8751 | 4.73 | 4.72 | 4.61 | 4.71 | 4.64 | 4.7085 | 4.8 | 4.91 | 4.8 | 4.03 | 4 | 5 | 4.8 | 5.5
| Mean |  | 4.9833 | 5.3333 | 5.5667 | 4.5333 | 10.5019 | 4.5517 | 3.65 | 4.65 | 4.9165 | 4.6017 | 4.6467 | 4.6333 | 4.6 | 4.665 | 4.7231 | 4.8367 | 4.675 | 4.8417 | 3.9517 | 4.3333 | 5 | 4.7233 | 5.5167
| Median |  | 4.975 | 5 | 5.5 | 4.5 | 10.5622 | 4.465 | 3.65 | 4.65 | 4.9252 | 4.585 | 4.66 | 4.635 | 4.595 | 4.65 | 4.7154 | 4.82 | 4.68 | 4.84 | 3.975 | 4 | 5 | 4.835 | 5.5
| Std Dev. |  | 0.14208 | 0.5164 | 0.17512 | 0.05164 | 0.12062 | 0.22085 | 0.10488 | 0.13784 | 0.18464 | 0.06969 | 0.069761 | 0.037771 | 0.078486 | 0.038341 | 0.11324 | 0.10211 | 0.15424 | 0.0392 | 0.067651 | 0.5164 | 0 | 0.36773 | 0.040825
| Rel.Std.Dev. |  | 2.851% | 9.682% | 3.146% | 1.139% | 1.149% | 4.852% | 2.873% | 2.964% | 3.755% | 1.514% | 1.501% | 0.8152% | 1.706% | 0.8219% | 2.398% | 2.111% | 3.299% | 0.8096% | 1.712% | 11.92% | 0% | 7.785% | 0.74%
| PDM3 |  | 5.906% | 13.34% | 18.3% | -3.657% | 123.2% | -3.268% | -22.43% | -1.178% | 4.487% | -2.205% | -1.249% | -1.532% | -2.24% | -0.859% | 0.3758% | 2.789% | -0.6465% | 2.896% | -16.02% | -7.908% | 6.26% | 0.3807% | 17.24%
| Z-Score (Absolute) |  | 1.66 | NA | 4.74 | 0.72 | 30.84 | 0.63 | 5.39 | 0.11 | 1.3 | 0.36 | 0.12 | 0.19 | 0.37 | 0.03 | 0.28 | 0.88 | 0.03 | 0.91 | 3.8 | NA | NA | 0.28 | 4.48
Results from laboratories 2, 20 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 553b (Certified Value 1.7 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.81 | 1.47 | 1.64 | 1.73 | 1.64 | 1.64 | 1.6918 | 1.53 | 1.71 | 1.64 | 1.66 | 1.74 | 1.7693 | 1.7 | 1.68 | 1.89 | NA | 1.6 | 1.7 | 1.71 | 1.23 |  |  | 1
| 1 | 2 | 1.79 | 1.43 | 1.63 | 1.73 | 1.69 | 1.57 | 1.8422 | 1.6 | 1.79 | 1.62 | 1.64 | 1.72 | 1.6878 | 1.68 | 1.71 | 1.9 | NA | 1.6 | 1.7 | 1.7 | 1.25 |  |  | 7
| 1 | 3 | 1.78 | 1.5 | 1.6 | 1.74 | 1.74 | 1.67 | 1.8865 | 1.6 | 1.72 | 1.62 | 1.56 | 1.73 | 1.7398 | 1.7 | 1.82 | 1.82 | NA | 1.7 | 1.7 | 1.65 | 1.27 |  |  | 16
| 1 | 4 | 1.81 | 1.47 | 1.63 | 1.79 | 1.69 | 1.62 | 1.6714 | 1.53 | 1.68 | 1.62 | 1.59 | 1.67 | 1.7524 | 1.7 | 1.73 | 1.83 | NA | 1.7 | 1.7 | 1.65 | 1.27 |  |  | 1.6997
| 1 | 5 | 1.8 | 1.58 | 1.65 | 1.72 | 1.65 | 1.58 | 1.6781 | 1.52 | 1.76 | 1.68 | 1.66 | 1.7 | 1.7264 | 1.7 | 1.76 | 1.88 | NA | 1.6 | 1.7 | 1.65 | 1.21 |  |  | 12
| 1 | 6 | 1.8 | 1.45 | 1.62 | 1.74 | 1.75 | 1.65 | 1.8752 | 1.62 | 1.77 | 1.66 | 1.67 | 1.66 | 1.7253 | 1.69 | 1.77 | 1.87 | NA | 1.7 | 1.7 | 1.66 | 1.13
| Mean |  | 1.7983 | 1.4833 | 1.6283 | 1.7417 | 1.6933 | 1.6217 | 1.7742 | 1.5667 | 1.7383 | 1.64 | 1.63 | 1.7033 | 1.7335 | 1.695 | 1.745 | 1.865 |  | 1.65 | 1.7 | 1.67 | 1.2267
| Median |  | 1.8 | 1.47 | 1.63 | 1.735 | 1.69 | 1.63 | 1.767 | 1.565 | 1.74 | 1.63 | 1.65 | 1.71 | 1.7331 | 1.7 | 1.745 | 1.875 |  | 1.65 | 1.7 | 1.655 | 1.24
| Std Dev. |  | 0.01169 | 0.052789 | 0.017224 | 0.024833 | 0.045019 | 0.039707 | 0.10393 | 0.044572 | 0.041673 | 0.025298 | 0.044721 | 0.03266 | 0.027852 | 0.0083666 | 0.049295 | 0.032711 |  | 0.054772 | 0 | 0.027568 | 0.052789
| Rel.Std.Dev. |  | 0.6501% | 3.559% | 1.058% | 1.426% | 2.659% | 2.449% | 5.858% | 2.845% | 2.397% | 1.543% | 2.744% | 1.917% | 1.607% | 0.4936% | 2.825% | 1.754% |  | 3.32% | 0% | 1.651% | 4.303%
| PDM3 |  | 5.804% | -12.73% | -4.198% | 2.47% | -0.3737% | -4.59% | 4.383% | -7.826% | 2.274% | -3.512% | -4.1% | 0.2146% | 1.989% | -0.2757% | 2.666% | 9.726% |  | -2.923% | 0.01848% | -1.747% | -27.83%
| Z-Score (Absolute) |  | 1.34 | 2.71 | 0.85 | 0.61 | 0.01 | 0.93 | 1.03 | 1.64 | 0.57 | 0.7 | 0.82 | 0.12 | 0.51 | 0.01 | 0.65 | 2.19 | NA | 0.57 | 0.07 | 0.31 | 6
| Analytical results for Ca in OREAS 553b (Certified Value 0.313 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.3063 | 0.29 | 0.32 | 0.314 | 0.31251 | 0.33 | 0.31 | 0.29 | 0.30878 | 0.32 | 0.31 | 0.32 | 0.33 | 0.32 | 0.32312 | 0.32 | 0.31 | 0.3206 | 0.3313 | 0.29 | 0.31 | 0.243 | 0.314 | 1
| 1 | 2 | 0.3069 | 0.29 | 0.31 | 0.306 | 0.31209 | 0.32 | 0.31 | 0.31 | 0.30213 | 0.32 | 0.32 | 0.32 | 0.32 | 0.31 | 0.32296 | 0.32 | 0.308 | 0.3217 | 0.3284 | 0.3 | 0.31 | 0.248 | 0.311 | #N/A
| 1 | 3 | 0.3087 | 0.3 | 0.31 | 0.302 | 0.31241 | 0.299 | 0.31 | 0.3 | 0.30721 | 0.32 | 0.3 | 0.32 | 0.33 | 0.31 | 0.32104 | 0.33 | 0.309 | 0.3252 | 0.3299 | 0.29 | 0.31 | 0.253 | 0.312 | 16
| 1 | 4 | 0.3131 | 0.28 | 0.33 | 0.303 | 0.31197 | 0.307 | 0.31 | 0.32 | 0.30234 | 0.31 | 0.3 | 0.32 | 0.33 | 0.31 | 0.32547 | 0.32 | 0.314 | 0.3237 | 0.3208 | 0.29 | 0.31 | 0.24 | 0.313 | 0.31286
| 1 | 5 | 0.3151 | 0.33 | 0.32 | 0.316 | 0.31186 | 0.307 | 0.31 | 0.29 | 0.30834 | 0.32 | 0.31 | 0.32 | 0.33 | 0.31 | 0.32328 | 0.32 | 0.312 | 0.3238 | 0.3208 | 0.29 | 0.31 | 0.247 | 0.306 | 13
| 1 | 6 | 0.3092 | 0.3 | 0.31 | 0.303 | 0.31132 | 0.279 | 0.3 | 0.31 | 0.30492 | 0.32 | 0.32 | 0.32 | 0.34 | 0.31 | 0.31496 | 0.32 | 0.311 | 0.326 | 0.3342 | 0.3 | 0.31 | 0.25 | 0.305
| Mean |  | 0.30988 | 0.29833 | 0.31667 | 0.30733 | 0.31203 | 0.307 | 0.30833 | 0.30333 | 0.30562 | 0.31833 | 0.31 | 0.32 | 0.33 | 0.31167 | 0.3218 | 0.32167 | 0.31067 | 0.3235 | 0.32757 | 0.29333 | 0.31 | 0.24683 | 0.31017
| Median |  | 0.30895 | 0.295 | 0.315 | 0.3045 | 0.31203 | 0.307 | 0.31 | 0.305 | 0.30607 | 0.32 | 0.31 | 0.32 | 0.33 | 0.31 | 0.32304 | 0.32 | 0.3105 | 0.32375 | 0.32915 | 0.29 | 0.31 | 0.2475 | 0.3115
| Std Dev. |  | 0.0034977 | 0.017224 | 0.008165 | 0.006121 | 0.00042777 | 0.017584 | 0.0040825 | 0.012111 | 0.0029456 | 0.0040825 | 0.0089443 | 0 | 0.0063246 | 0.0040825 | 0.0036373 | 0.0040825 | 0.0021602 | 0.0020455 | 0.0055795 | 0.005164 | 0 | 0.0047081 | 0.0037639
| Rel.Std.Dev. |  | 1.129% | 5.773% | 2.578% | 1.992% | 0.1371% | 5.728% | 1.324% | 3.993% | 0.9638% | 1.282% | 2.885% | 0% | 1.917% | 1.31% | 1.13% | 1.269% | 0.6954% | 0.6323% | 1.703% | 1.76% | 0% | 1.907% | 1.213%
| PDM3 |  | -0.9502% | -4.642% | 1.218% | -1.765% | -0.2651% | -1.872% | -1.446% | -3.044% | -2.313% | 1.751% | -0.9129% | 2.283% | 5.48% | -0.3802% | 2.86% | 2.816% | -0.6998% | 3.402% | 4.702% | -6.24% | -0.9129% | -21.1% | -0.8597%
| Z-Score (Absolute) |  | 0.03 | 1.23 | 0.67 | 0.29 | 0.19 | 0.33 | 0.19 | 0.71 | 0.47 | 0.85 | 0.02 | 1.02 | 2.06 | 0.16 | 1.21 | 1.19 | 0.05 | 1.38 | 1.81 | 1.75 | 0.02 | 6.57 | 0
| Analytical results for Cd in OREAS 553b (Certified Value 0.3 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.29 | 0.45 | < 0.3 | 0.3 | 0.31 | < 0.5 | 0.3 | 0.30954 | 0.29 | 0.3 | 0.29 | 0.31 | 0.29 | < 0.5 | 0.31 | 0.29 | 0.33 | 0.2 | 0.3 | 0.24 | 0.3 |  |  | 1
| 1 | 2 | 0.3 | 0.43 | < 0.3 | 0.31 | 0.3 | < 0.5 | 0.28 | 0.29594 | 0.29 | 0.28 | 0.29 | 0.3 | 0.26 | < 0.5 | 0.32 | 0.3 | 0.32 | 0.3 | 0.3 | 0.23 | 0.3 |  |  | 8
| 1 | 3 | 0.32 | 0.48 | < 0.3 | 0.3 | 0.29 | < 0.5 | 0.29 | 0.32809 | 0.31 | 0.29 | 0.28 | 0.27 | 0.29 | < 0.5 | 0.34 | 0.3 | 0.34 | 0.3 | 0.5 | 0.26 | 0.28 |  |  | 16
| 1 | 4 | 0.32 | 0.43 | < 0.3 | 0.29 | 0.28 | < 0.5 | 0.32 | 0.33488 | 0.29 | 0.31 | 0.26 | 0.29 | 0.27 | < 0.5 | 0.36 | 0.3 | 0.31 | 0.3 | 0.3 | 0.22 | 0.32 |  |  | 0.29591
| 1 | 5 | 0.33 | 0.45 | < 0.3 | 0.29 | 0.31 | < 0.5 | 0.28 | 0.28961 | 0.31 | 0.3 | 0.28 | 0.31 | 0.27 | < 0.5 | 0.34 | 0.3 | 0.31 | 0.2 | 0.3 | 0.23 | 0.28 |  |  | 14
| 1 | 6 | 0.29 | 0.41 | < 0.3 | 0.27 | 0.28 | < 0.5 | 0.31 | 0.3172 | 0.29 | 0.3 | 0.28 | 0.32 | 0.28 | < 0.5 | 0.34 | 0.31 | 0.35 | 0.3 | 0.3 | 0.25 | 0.29
| Mean |  | 0.30833 | 0.44167 |  | 0.29333 | 0.295 |  | 0.29667 | 0.31254 | 0.29667 | 0.29667 | 0.28 | 0.3 | 0.27667 |  | 0.335 | 0.3 | 0.32667 | 0.26667 | 0.33333 | 0.23833 | 0.295
| Median |  | 0.31 | 0.44 |  | 0.295 | 0.295 |  | 0.295 | 0.31337 | 0.29 | 0.3 | 0.28 | 0.305 | 0.275 |  | 0.34 | 0.3 | 0.325 | 0.3 | 0.3 | 0.235 | 0.295
| Std Dev. |  | 0.017224 | 0.024014 |  | 0.013663 | 0.013784 |  | 0.01633 | 0.017737 | 0.010328 | 0.010328 | 0.010954 | 0.017889 | 0.012111 |  | 0.017607 | 0.0063246 | 0.01633 | 0.05164 | 0.08165 | 0.01472 | 0.015166
| Rel.Std.Dev. |  | 5.586% | 5.437% |  | 4.658% | 4.673% |  | 5.504% | 5.675% | 3.481% | 3.481% | 3.912% | 5.963% | 4.377% |  | 5.256% | 2.108% | 4.999% | 19.36% | 24.49% | 6.176% | 5.141%
| PDM3 |  | 4.2% | 49.26% |  | -0.8695% | -0.3062% |  | 0.257% | 5.622% | 0.257% | 0.257% | -5.375% | 1.383% | -6.502% |  | 13.21% | 1.383% | 10.4% | -9.881% | 12.65% | -19.46% | -0.3062%
| Z-Score (Absolute) |  | 0.5 | 6.16 | 6.23 | 0.14 | 0.07 | 1.98 | 0 | 0.67 | 0 | 0 | 0.71 | 0.14 | 0.85 | 1.98 | 1.63 | 0.14 | 1.27 | NA | NA | 2.48 | 0.07
Results from laboratories 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 553b (Certified Value 144 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
0
| 1 | 1 | 140.4 | 131.88 | 176 | 184.096 | 148 | 124 | 117 | 156.5 | 139 | 139.5 | 149 | 153.5 | 147.176 | 153 | 128 | 149.33 | 162.2 | 99 | 100 | 149 | 112 |  |  | 1
| 1 | 2 | 141.03 | 131.56 | 179 | 184.176 | 136 | 121 | 112.17 | 154 | 149 | 137.5 | 146 | 147 | 144.4459 | 148 | 131 | 151.76 | 159.8 | 106 | 112 | 148 | 114 |  |  | 9
| 1 | 3 | 141.2 | 129.68 | 173 | 184.184 | 139 | 126 | 116.39 | 156.5 | 139.5 | 137.5 | 145 | 151.5 | 151.0536 | 150 | 131 | 148.67 | 162 | 102 | 105 | 147 | 108 |  |  | 16
| 1 | 4 | 144.46 | 129.46 | 180 | 184.176 | 138 | 123 | 113.97 | 154.5 | 122.5 | 135 | 147 | 150.5 | 147.4773 | 152 | 136 | 150.57 | 156 | 112 | 111 | 164 | 112 |  |  | 144.2478
| 1 | 5 | 145.93 | 144.48 | 178 | 184.16 | 144 | 118 | 111.34 | 151.5 | 142.5 | 136 | 140 | 148 | 147.2892 | 152 | 144 | 151.75 | 159.8 | 111 | 105 | 145 | 99 |  |  | 15
| 1 | 6 | 147.18 | 133.35 | 176 | 184.16 | 133 | 124 | 115.88 | 163.5 | 140 | 137 | 150 | 146 | 146.473 | 152 | 135 | 155.72 | 163.4 | 123 | 103 | 145 | 88
| Mean |  | 143.3667 | 133.4017 | 177 | 184.1587 | 139.6667 | 122.6667 | 114.4583 | 156.0833 | 138.75 | 137.0833 | 146.1667 | 149.4167 | 147.3192 | 151.1667 | 134.1667 | 151.3 | 160.5333 | 108.8333 | 106 | 149.6667 | 105.5
| Median |  | 142.83 | 131.72 | 177 | 184.168 | 138.5 | 123.5 | 114.925 | 155.5 | 139.75 | 137.25 | 146.5 | 149.25 | 147.2326 | 152 | 133 | 151.16 | 160.9 | 108.5 | 105 | 147.5 | 110
| Std Dev. |  | 2.8727 | 5.6183 | 2.5298 | 0.032166 | 5.465 | 2.8048 | 2.3414 | 4.0794 | 8.7792 | 1.5303 | 3.5449 | 2.8882 | 2.1431 | 1.8348 | 5.6362 | 2.5012 | 2.6372 | 8.5654 | 4.6476 | 7.2019 | 10.1143
| Rel.Std.Dev. |  | 2.004% | 4.212% | 1.429% | 0.01747% | 3.913% | 2.286% | 2.046% | 2.614% | 6.327% | 1.116% | 2.425% | 1.933% | 1.455% | 1.214% | 4.201% | 1.653% | 1.643% | 7.87% | 4.385% | 4.812% | 9.587%
| PDM3 |  | -0.6108% | -7.519% | 22.71% | 27.67% | -3.176% | -14.96% | -20.65% | 8.205% | -3.811% | -4.967% | 1.33% | 3.583% | 2.129% | 4.797% | -6.989% | 4.889% | 11.29% | -24.55% | -26.52% | 3.757% | -26.86%
| Z-Score (Absolute) |  | 0 | 0.73 | 2.47 | 2.99 | 0.27 | 1.52 | 2.12 | 0.93 | 0.34 | 0.46 | 0.21 | 0.44 | 0.29 | 0.57 | 0.67 | 0.58 | 1.26 | 2.53 | 2.74 | 0.46 | 2.78
| Analytical results for Co in OREAS 553b (Certified Value 0.587 wt.%) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*MS | 4A*AAS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.58154 | 0.56 | 0.61 | 0.598 | 0.5167 | 0.3678 | 0.596 | 0.6022 | 0.58833 | 0.561 | 0.589 | 0.575 | 0.575 | 0.59 | 0.58985 | 0.5855 | 0.5837 | 0.54662 | 0.60125 | 0.602 | 0.5736 | 0.5835 | 0.5973 | 1
| 1 | 2 | 0.57784 | 0.58 | 0.597 | 0.609 | 0.51612 | 0.3652 | 0.5959 | 0.5883 | 0.57993 | 0.554 | 0.593 | 0.569 | 0.577 | 0.583 | 0.58513 | 0.59 | 0.5899 | 0.54967 | 0.60046 | 0.606 | 0.58416 | 0.58907 | 0.582 | 30
| 1 | 3 | 0.58381 | 0.58 | 0.594 | 0.599 | 0.51649 | 0.3608 | 0.5935 | 0.5948 | 0.58692 | 0.563 | 0.566 | 0.573 | 0.572 | 0.583 | 0.59142 | 0.5942 | 0.5811 | 0.54396 | 0.59803 | 0.607 | 0.57319 | 0.58893 | 0.5835 | 16
| 1 | 4 | 0.58624 | 0.58 | 0.597 | 0.603 | 0.51646 | 0.3655 | 0.5954 | 0.5908 | 0.58732 | 0.554 | 0.575 | 0.582 | 0.579 | 0.578 | 0.57804 | 0.5943 | 0.5851 | 0.54449 | 0.5821 | 0.611 | 0.58314 | 0.58524 | 0.5852 | 0.58712
| 1 | 5 | 0.59659 | 0.58 | 0.603 | 0.603 | 0.51606 | 0.3785 | 0.5955 | 0.6048 | 0.58015 | 0.556 | 0.572 | 0.574 | 0.596 | 0.584 | 0.58688 | 0.5943 | 0.5867 | 0.55423 | 0.58893 | 0.608 | 0.58957 | 0.58523 | 0.5738 | 16
| 1 | 6 | 0.59667 | 0.58 | 0.604 | 0.602 | 0.51646 | 0.3579 | 0.5969 | 0.5952 | 0.58743 | 0.565 | 0.591 | 0.573 | 0.596 | 0.575 | 0.58211 | 0.5905 | 0.5978 | 0.55169 | 0.61058 | 0.607 | 0.58959 | 0.58319 | 0.5815
| Mean |  | 0.58711 | 0.57667 | 0.60083 | 0.60233 | 0.51638 | 0.36595 | 0.59553 | 0.59602 | 0.58501 | 0.55883 | 0.581 | 0.57433 | 0.5825 | 0.58217 | 0.58557 | 0.59147 | 0.58738 | 0.54844 | 0.59689 | 0.60683 | 0.58221 | 0.58586 | 0.58388
| Median |  | 0.58503 | 0.58 | 0.6 | 0.6025 | 0.51646 | 0.36535 | 0.5957 | 0.595 | 0.58712 | 0.5585 | 0.582 | 0.5735 | 0.578 | 0.583 | 0.586 | 0.59235 | 0.5859 | 0.54814 | 0.59925 | 0.607 | 0.58365 | 0.58524 | 0.58275
| Std Dev. |  | 0.0078723 | 0.008165 | 0.0059133 | 0.0038816 | 0.00024338 | 0.0071136 | 0.001129 | 0.006392 | 0.003879 | 0.0047924 | 0.011402 | 0.004274 | 0.01071 | 0.0051929 | 0.0049632 | 0.0035274 | 0.0058918 | 0.004113 | 0.010026 | 0.0029269 | 0.007332 | 0.0025771 | 0.0076523
| Rel.Std.Dev. |  | 1.341% | 1.416% | 0.9842% | 0.6444% | 0.04713% | 1.944% | 0.1896% | 1.072% | 0.6631% | 0.8576% | 1.962% | 0.7442% | 1.839% | 0.892% | 0.8476% | 0.5964% | 1.003% | 0.7499% | 1.68% | 0.4823% | 1.259% | 0.4399% | 1.311%
| PDM3 |  | -0.001191% | -1.781% | 2.335% | 2.591% | -12.05% | -37.67% | 1.433% | 1.515% | -0.3594% | -4.818% | -1.043% | -2.178% | -0.7872% | -0.844% | -0.2643% | 0.74% | 0.04451% | -6.588% | 1.664% | 3.357% | -0.8369% | -0.2151% | -0.5516%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 1.28 | 1.4 | 5.55 | 17.7 | 0.85 | 0.89 | 0 | 2.12 | 0.32 | 0.86 | 0.2 | 0.23 | 0.05 | 0.52 | 0.19 | 2.95 | 0.96 | 1.76 | 0.23 | 0.07 | 0.09
| Analytical results for Cr in OREAS 553b (Certified Value 91 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 94 | 110 | 44 | 88 | 107 | 89 | 88 | 91.7603 | 97 | 102 | 85 | 88 | 84 | 100.7426 | 76.1 | 97 | 104 | 84.8 | 72 | 88 | 72 | 88 |  | 1
| 1 | 2 | 96 | 109 | 53 | 85 | 112 | 88 | 82 | 94.6974 | 97 | 103 | 85 | 84 | 86 | 98.7914 | 75.6 | 96 | 101 | 82.7 | 73 | 87 | 77 | 86 |  | #N/A
| 1 | 3 | 97 | 108 | 56 | 84 | 106 | 91 | 85 | 92.8959 | 98 | 99 | 82 | 80 | 82 | 101.3265 | 76.8 | 97 | 103 | 82.7 | 70 | 86 | 74 | 81 |  | 16
| 1 | 4 | 98 | 108 | 63 | 89 | 107 | 89 | 87 | 93.2683 | 97 | 100 | 84 | 86 | 88 | 100.6685 | 75.9 | 98 | 100 | 80.4 | 72 | 86 | 76 | 88 |  | 90.5935
| 1 | 5 | 98 | 112 | 60 | 87 | 101 | 88 | 81 | 92.5113 | 98 | 100 | 88 | 86 | 91 | 100.4262 | 75.3 | 99 | 101 | 81.9 | 69 | 86 | 73 | 98 |  | 17
| 1 | 6 | 97 | 108 | 51 | 86 | 100 | 91 | 86 | 94.6015 | 99 | 102 | 87 | 86 | 89 | 98.4905 | 75.6 | 97 | 104 | 84.8 | 73 | 86 | 74 | 105
| Mean |  | 96.6667 | 109.1667 | 54.5 | 86.5 | 105.5 | 89.3333 | 84.8333 | 93.2891 | 97.6667 | 101 | 85.1667 | 85 | 86.6667 | 100.0743 | 75.8833 | 97.3333 | 102.1667 | 82.8833 | 71.5 | 86.5 | 74.3333 | 91
| Median |  | 97 | 108.5 | 54.5 | 86.5 | 106.5 | 89 | 85.5 | 93.0821 | 97.5 | 101 | 85 | 86 | 87 | 100.5473 | 75.75 | 97 | 102 | 82.7 | 72 | 86 | 74 | 88
| Std Dev. |  | 1.5055 | 1.6021 | 6.775 | 1.8708 | 4.4159 | 1.3663 | 2.7869 | 1.1664 | 0.8165 | 1.5492 | 2.137 | 2.7568 | 3.3267 | 1.1529 | 0.52694 | 1.0328 | 1.7224 | 1.7058 | 1.6432 | 0.83666 | 1.8619 | 8.8091
| Rel.Std.Dev. |  | 1.557% | 1.468% | 12.43% | 2.163% | 4.186% | 1.529% | 3.285% | 1.25% | 0.836% | 1.534% | 2.509% | 3.243% | 3.838% | 1.152% | 0.6944% | 1.061% | 1.686% | 2.058% | 2.298% | 0.9672% | 2.505% | 9.68%
| PDM3 |  | 6.704% | 20.5% | -39.84% | -4.519% | 16.45% | -1.391% | -6.358% | 2.976% | 7.808% | 11.49% | -5.99% | -6.174% | -4.335% | 10.47% | -16.24% | 7.44% | 12.77% | -8.511% | -21.08% | -4.519% | -17.95% | 0.4487%
| Z-Score (Absolute) |  | 0.65 | 1.59 | 2.51 | 0.11 | 1.31 | 0.1 | 0.24 | 0.4 | 0.72 | 0.97 | 0.21 | 0.22 | 0.1 | 0.9 | 0.91 | 0.7 | 1.06 | 0.38 | 1.24 | 0.11 | 1.02 | 0.22
| Analytical results for Cs in OREAS 553b (Certified Value 2.08 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.1 | 1.9 | 2.19 | 2.1 | 2.752 | 2.07 | 2.31 | 2.23 | 2.1328 | 2.08 | 1.97 | 1.95 | 2.11 | 2.08 | 2.1674 | 1.98 | 2.07 | 2.22 | NA | 2 | 2.1 | 1.58 |  | 1
| 1 | 2 | 2.12 | 1.7 | 2.15 | 2.1 | 2.6818 | 1.97 | 2.15 | 2.09 | 2.2246 | 2.01 | 2.14 | 1.9 | 2.03 | 2 | 2.1286 | 2.01 | 2.1 | 2.2 | NA | 2.4 | 2.1 | 1.57 |  | 12
| 1 | 3 | 2.15 | 1.8 | 2.19 | 2.1 | 2.7201 | 2.01 | 2.21 | 2.25 | 2.2641 | 2.05 | 2 | 1.92 | 2.08 | 2.01 | 2.1779 | 1.95 | 2.02 | 2.17 | NA | 2 | 2.1 | 1.53 |  | 16
| 1 | 4 | 2.16 | 1.9 | 2.24 | 2.1 | 2.7552 | 1.95 | 2.1 | 2.07 | 2.0616 | 2 | 1.92 | 1.92 | 2.1 | 1.99 | 2.1524 | 2.01 | 2.07 | 2.22 | NA | 2.2 | 2 | 1.53 |  | 2.0815
| 1 | 5 | 2.19 | 2.1 | 2.16 | 2.1 | 2.8874 | 2.08 | 2.02 | 2.17 | 2.0211 | 1.97 | 2.09 | 1.9 | 2.02 | 2 | 2.06 | 1.93 | 2.11 | 2.23 | NA | 2.2 | 2 | 1.5 |  | 18
| 1 | 6 | 2.22 | 1.9 | 2.12 | 2.1 | 2.9926 | 1.94 | 2.33 | 2.2 | 2.1211 | 2.11 | 2.08 | 1.94 | 2.06 | 1.98 | 2.1102 | 2 | 2.11 | 2.2 | NA | 2.2 | 2.1 | 1.52
| Mean |  | 2.1567 | 1.8833 | 2.175 | 2.1 | 2.7982 | 2.0033 | 2.1867 | 2.1683 | 2.1375 | 2.0367 | 2.0333 | 1.9217 | 2.0667 | 2.01 | 2.1328 | 1.98 | 2.08 | 2.2067 |  | 2.1667 | 2.0667 | 1.5383
| Median |  | 2.155 | 1.9 | 2.175 | 2.1 | 2.7536 | 1.99 | 2.18 | 2.185 | 2.1269 | 2.03 | 2.04 | 1.92 | 2.07 | 2 | 2.1405 | 1.99 | 2.085 | 2.21 |  | 2.2 | 2.1 | 1.53
| Std Dev. |  | 0.044121 | 0.13292 | 0.041352 | 0 | 0.1178 | 0.060553 | 0.12078 | 0.073869 | 0.092982 | 0.052789 | 0.083347 | 0.020412 | 0.036697 | 0.035777 | 0.043485 | 0.033466 | 0.034641 | 0.021602 |  | 0.15055 | 0.05164 | 0.030605
| Rel.Std.Dev. |  | 2.046% | 7.057% | 1.901% | 0% | 4.21% | 3.023% | 5.523% | 3.407% | 4.35% | 2.592% | 4.099% | 1.062% | 1.776% | 1.78% | 2.039% | 1.69% | 1.665% | 0.979% |  | 6.949% | 2.499% | 1.989%
| PDM3 |  | 3.611% | -9.521% | 4.492% | 0.8885% | 34.43% | -3.756% | 5.052% | 4.171% | 2.692% | -2.154% | -2.314% | -7.679% | -0.713% | -3.435% | 2.462% | -4.877% | -0.07239% | 6.013% |  | 4.091% | -0.713% | -26.1%
| Z-Score (Absolute) |  | 0.67 | 1.73 | 0.84 | 0.18 | 6.32 | 0.67 | 0.94 | 0.78 | 0.51 | 0.38 | 0.41 | 1.39 | 0.12 | 0.62 | 0.46 | 0.88 | 0 | 1.11 | NA | 0.76 | 0.12 | 4.76
| Analytical results for Cu in OREAS 553b (Certified Value 3.6 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 3.6037 | 3.455 | 3.6 | 3.57 | 3.5457 | > 1 | 3.65 | 3.671 | 3.5702 | 3.52 | 3.58 | 3.57 | 3.6 | 3.56 | 3.7204 | 3.622 | 3.6 | 3.5934 | 3.6808 | 3.606 | 3.68 | 3.712 | 3.501 | 1
| 1 | 2 | 3.6894 | 3.605 | 3.58 | 3.54 | 3.5448 | > 1 | 3.65 | 3.733 | 3.5795 | 3.56 | 3.48 | 3.57 | 3.6 | 3.55 | 3.6864 | 3.638 | 3.62 | 3.5558 | 3.6145 | 3.599 | 3.67 | 3.642 | 3.5146 | 31
| 1 | 3 | 3.6605 | 3.574 | 3.53 | 3.53 | 3.5485 | > 1 | 3.65 | 3.69 | 3.5846 | 3.57 | 3.59 | 3.56 | 3.59 | 3.58 | 3.7502 | 3.63 | 3.53 | 3.6838 | 3.6326 | 3.596 | 3.66 | 3.667 | 3.5444 | 16
| 1 | 4 | 3.6868 | 3.635 | 3.53 | 3.57 | 3.5244 | > 1 | 3.63 | 3.673 | 3.6332 | 3.55 | 3.62 | 3.57 | 3.59 | 3.57 | 3.6876 | 3.613 | 3.66 | 3.6171 | 3.5298 | 3.637 | 3.62 | 3.661 | 3.5148 | 3.6027
| 1 | 5 | 3.635 | 3.544 | 3.53 | 3.46 | 3.553 | > 1 | 3.64 | 3.709 | 3.655 | 3.55 | 3.53 | 3.55 | 3.57 | 3.54 | 3.7069 | 3.639 | 3.65 | 3.6228 | 3.5512 | 3.626 | 3.65 | 3.589 | 3.5249 | 19
| 1 | 6 | 3.6906 | 3.632 | 3.52 | 3.56 | 3.545 | > 1 | 3.65 | 3.736 | 3.669 | 3.59 | 3.47 | 3.57 | 3.62 | 3.55 | 3.6786 | 3.625 | 3.62 | 3.6011 | 3.6403 | 3.661 | 3.59 | 3.696 | 3.5137
| Mean |  | 3.661 | 3.5742 | 3.5483 | 3.5383 | 3.5436 |  | 3.645 | 3.702 | 3.6153 | 3.5567 | 3.545 | 3.565 | 3.595 | 3.5583 | 3.705 | 3.6278 | 3.6133 | 3.6123 | 3.6082 | 3.6208 | 3.645 | 3.6612 | 3.5189
| Median |  | 3.6736 | 3.5895 | 3.53 | 3.55 | 3.5454 |  | 3.65 | 3.6995 | 3.6089 | 3.555 | 3.555 | 3.57 | 3.595 | 3.555 | 3.6972 | 3.6275 | 3.62 | 3.6091 | 3.6236 | 3.616 | 3.655 | 3.664 | 3.5147
| Std Dev. |  | 0.035519 | 0.067981 | 0.033116 | 0.041673 | 0.0098875 |  | 0.0083666 | 0.028678 | 0.042514 | 0.023381 | 0.061563 | 0.0083666 | 0.016432 | 0.01472 | 0.026954 | 0.0099482 | 0.046332 | 0.042238 | 0.057143 | 0.025357 | 0.033912 | 0.043375 | 0.014622
| Rel.Std.Dev. |  | 0.9702% | 1.902% | 0.9333% | 1.178% | 0.279% |  | 0.2295% | 0.7746% | 1.176% | 0.6574% | 1.737% | 0.2347% | 0.4571% | 0.4137% | 0.7275% | 0.2742% | 1.282% | 1.169% | 1.584% | 0.7003% | 0.9304% | 1.185% | 0.4155%
| PDM3 |  | 1.617% | -0.7932% | -1.51% | -1.788% | -1.643% |  | 1.173% | 2.755% | 0.348% | -1.279% | -1.603% | -1.048% | -0.2149% | -1.233% | 2.839% | 0.6964% | 0.294% | 0.2662% | 0.1523% | 0.5021% | 1.173% | 1.622% | -2.327%
| Z-Score (Absolute) |  | 0.71 | 0.51 | 0.87 | 1.01 | 0.94 | NA | 0.49 | 1.29 | 0.07 | 0.75 | 0.92 | 0.64 | 0.21 | 0.73 | 1.33 | 0.25 | 0.04 | 0.03 | 0.03 | 0.15 | 0.49 | 0.71 | 1.28
| Analytical results for Dy in OREAS 553b (Certified Value 4.44 ppm) | [24 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 4.71 | 3.9 | 5.3 | 5.5386 | 4.7 | 4.7501 | 4.34 | 4.51 | [15 empty] | 1
| 1 | 2 | 4.3 | 3.7 | 5.8 | 5.6452 | 4.49 | 4.3838 | 4.44 | 4.65 | [15 empty] | 13
| 1 | 3 | 4.5 | 3.9 | 5.6 | 5.586 | 4.76 | 4.9247 | 4.48 | 4.65 | [15 empty] | 16
| 1 | 4 | 4.45 | 3.9 | 5.8 | 5.5498 | 4.69 | 4.7636 | 4.75 | 4.73 | [15 empty] | 4.4419
| 1 | 5 | 4.6 | 4 | 5.7 | 5.5861 | 4.76 | 4.5985 | 4.12 | 4.31 | [15 empty] | 20
| 1 | 6 | 4.51 | 3.9 | 5.6 | 5.5275 | 4.52 | 4.6692 | 4.58 | 3.97
| Mean |  | 4.5117 | 3.8833 | 5.6333 | 5.5722 | 4.6533 | 4.6816 | 4.4517 | 4.47
| Median |  | 4.505 | 3.9 | 5.65 | 5.5679 | 4.695 | 4.7097 | 4.46 | 4.58
| Std Dev. |  | 0.13848 | 0.098319 | 0.18619 | 0.043208 | 0.11894 | 0.18235 | 0.21395 | 0.28622
| Rel.Std.Dev. |  | 3.069% | 2.532% | 3.305% | 0.7754% | 2.556% | 3.895% | 4.806% | 6.403%
| PDM3 |  | 1.57% | -12.58% | 26.82% | 25.45% | 4.759% | 5.396% | 0.219% | 0.6317%
| Z-Score (Absolute) |  | 0.39 | 3.87 | 5.82 | 5.49 | 0.39 | 0.55 | 0.73 | 0.62
| Analytical results for Er in OREAS 553b (Certified Value 2.85 ppm) | [24 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 2.62 | 2.4 | 3.1 | 3.6473 | 3.2 | 2.9215 | 2.73 | 2.81 | [15 empty] | 1
| 1 | 2 | 2.68 | 2.3 | 3.4 | 3.6602 | 2.93 | 2.8603 | 2.86 | 2.88 | [15 empty] | 14
| 1 | 3 | 2.7 | 2.4 | 3.2 | 3.6582 | 3.12 | 3.0141 | 2.7 | 2.85 | [15 empty] | 16
| 1 | 4 | 2.74 | 2.4 | 3.4 | 3.6937 | 2.99 | 3.1006 | 2.92 | 2.92 | [15 empty] | 2.8486
| 1 | 5 | 2.71 | 2.6 | 3.3 | 3.6112 | 3.19 | 2.8908 | 2.64 | 2.75 | [15 empty] | 21
| 1 | 6 | 2.66 | 2.4 | 3.3 | 3.6851 | 2.89 | 2.9549 | 2.72 | 2.49
| Mean |  | 2.685 | 2.4167 | 3.2833 | 3.6593 | 3.0533 | 2.957 | 2.7617 | 2.7833
| Median |  | 2.69 | 2.4 | 3.3 | 3.6592 | 3.055 | 2.9382 | 2.725 | 2.83
| Std Dev. |  | 0.041833 | 0.098319 | 0.1169 | 0.029356 | 0.13456 | 0.088167 | 0.10591 | 0.15513
| Rel.Std.Dev. |  | 1.558% | 4.068% | 3.561% | 0.8022% | 4.407% | 2.982% | 3.835% | 5.574%
| PDM3 |  | -5.744% | -15.16% | 15.26% | 28.46% | 7.186% | 3.806% | -3.053% | -2.292%
| Z-Score (Absolute) |  | 0.68 | 1.66 | 1.51 | 2.89 | 0.67 | 0.32 | 0.4 | 0.32
| Analytical results for Eu in OREAS 553b (Certified Value 0.81 ppm) | [24 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.86 | 0.8 | 0.94 | 1.2374 | 0.78 | 0.93299 | 0.87 | 0.61 | [15 empty] | 1
| 1 | 2 | 0.89 | 0.7 | 0.93 | 1.2386 | 0.68 | 0.88415 | 0.86 | 0.6 | [15 empty] | 15
| 1 | 3 | 0.9 | 0.8 | 0.92 | 1.2484 | 0.72 | 0.9148 | 0.84 | 0.59 | [15 empty] | 16
| 1 | 4 | 0.85 | 0.8 | 0.95 | 1.2665 | 0.72 | 0.89748 | 0.96 | 0.61 | [15 empty] | 0.80926
| 1 | 5 | 0.87 | 0.8 | 0.94 | 1.2802 | 0.74 | 0.86999 | 0.83 | 0.57 | [15 empty] | 22
| 1 | 6 | 0.89 | 0.8 | 0.93 | 1.2555 | 0.67 | 0.85371 | 0.82 | 0.47
| Mean |  | 0.87667 | 0.78333 | 0.935 | 1.2544 | 0.71833 | 0.89219 | 0.86333 | 0.575
| Median |  | 0.88 | 0.8 | 0.935 | 1.252 | 0.72 | 0.89081 | 0.85 | 0.595
| Std Dev. |  | 0.019664 | 0.040825 | 0.010488 | 0.016646 | 0.040208 | 0.029129 | 0.050859 | 0.053572
| Rel.Std.Dev. |  | 2.243% | 5.212% | 1.122% | 1.327% | 5.597% | 3.265% | 5.891% | 9.317%
| PDM3 |  | 8.329% | -3.204% | 15.54% | 55.01% | -11.24% | 10.25% | 6.681% | -28.95%
| Z-Score (Absolute) |  | 0.06 | 0.77 | 0.58 | 3.42 | 1.35 | 0.2 | 0.06 | 2.62
| Analytical results for Fe in OREAS 553b (Certified Value 3.18 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 3.22 | 3.19 | 3.27 | 3.41 | 3.1349 | 3.3 | 3.13 | 3.16 | 3.2131 | 3.14 | 3.18 | 3.15 | 3.26 | 3.18 | 3.2609 | 3.19 | 3.24 | 3.33 | 3.2392 | 3.06 | 3.19 | 2.75 | 3.27 | 1
| 1 | 2 | 3.21 | 3.14 | 3.21 | 3.39 | 3.1322 | 3.15 | 3.13 | 3.2 | 3.1728 | 3.09 | 3.22 | 3.11 | 3.22 | 3.13 | 3.2707 | 3.15 | 3.25 | 3.34 | 3.2345 | 3.07 | 3.16 | 2.81 | 3.26 | #N/A
| 1 | 3 | 3.25 | 3.13 | 3.18 | 3.26 | 3.1257 | 2.96 | 3.16 | 3.13 | 3.1979 | 3.15 | 3.08 | 3.13 | 3.22 | 3.14 | 3.2659 | 3.18 | 3.24 | 3.29 | 3.2007 | 3.04 | 3.14 | 2.89 | 3.24 | 16
| 1 | 4 | 3.29 | 3.1 | 3.26 | 3.25 | 3.1281 | 3.09 | 3.13 | 3.22 | 3.1765 | 3.09 | 3.09 | 3.16 | 3.26 | 3.12 | 3.2889 | 3.19 | 3.28 | 3.32 | 3.1395 | 3.03 | 3.15 | 2.76 | 3.28 | 3.1829
| 1 | 5 | 3.34 | 3.29 | 3.29 | 3.41 | 3.1305 | 3.06 | 3.1 | 3.18 | 3.1624 | 3.09 | 3.09 | 3.15 | 3.34 | 3.15 | 3.2611 | 3.2 | 3.27 | 3.31 | 3.1631 | 3.02 | 3.15 | 2.79 | 3.15 | 23
| 1 | 6 | 3.22 | 3.12 | 3.19 | 3.39 | 3.1279 | 2.77 | 3.08 | 3.12 | 3.1972 | 3.17 | 3.2 | 3.13 | 3.36 | 3.11 | 3.2039 | 3.19 | 3.24 | 3.33 | 3.2896 | 3.09 | 3.14 | 2.89 | 3.13
| Mean |  | 3.255 | 3.1617 | 3.2333 | 3.3517 | 3.1299 | 3.055 | 3.1217 | 3.1683 | 3.1867 | 3.1217 | 3.1433 | 3.1383 | 3.2767 | 3.1383 | 3.2586 | 3.1833 | 3.2533 | 3.32 | 3.2111 | 3.0517 | 3.155 | 2.815 | 3.2217
| Median |  | 3.235 | 3.135 | 3.235 | 3.39 | 3.1293 | 3.075 | 3.13 | 3.17 | 3.1869 | 3.115 | 3.135 | 3.14 | 3.26 | 3.135 | 3.2635 | 3.19 | 3.245 | 3.325 | 3.2176 | 3.05 | 3.15 | 2.8 | 3.25
| Std Dev. |  | 0.050892 | 0.06969 | 0.045898 | 0.075476 | 0.0033367 | 0.17919 | 0.027869 | 0.0392 | 0.019054 | 0.036009 | 0.063456 | 0.018348 | 0.059889 | 0.024833 | 0.028714 | 0.017512 | 0.017512 | 0.017889 | 0.054842 | 0.026394 | 0.018708 | 0.061887 | 0.064936
| Rel.Std.Dev. |  | 1.564% | 2.204% | 1.42% | 2.252% | 0.1066% | 5.866% | 0.8928% | 1.237% | 0.5979% | 1.154% | 2.019% | 0.5847% | 1.828% | 0.7913% | 0.8812% | 0.5501% | 0.5383% | 0.5388% | 1.708% | 0.8649% | 0.593% | 2.198% | 2.016%
| PDM3 |  | 2.264% | -0.6681% | 1.584% | 5.301% | -1.667% | -4.019% | -1.925% | -0.4586% | 0.117% | -1.925% | -1.244% | -1.401% | 2.945% | -1.401% | 2.377% | 0.01266% | 2.212% | 4.306% | 0.885% | -4.124% | -0.8775% | -11.56% | 1.217%
| Z-Score (Absolute) |  | 1.25 | 0.1 | 0.94 | 2.65 | 0.56 | 1.64 | 0.67 | 0 | 0.26 | 0.67 | 0.36 | 0.43 | 1.57 | 0.43 | 1.3 | 0.22 | 1.23 | 2.19 | 0.62 | 1.69 | 0.19 | 5.11 | 0.77
| Analytical results for Ga in OREAS 553b (Certified Value 27.9 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 26.72 | 29.3 | 28.7 | 27.4 | 28 | 27.9 | 26.8563 | 27.4 | 28.2 | 28.4 | 25.9 | 28.7 | 28.9447 | 29.18 | 26.79 | 28.99 | 22.75 | 29 | 25.8 | 24.58 |  |  |  | 1
| 1 | 2 | 26.63 | 28.56 | 28 | 27.5 | 27.5 | 27.1 | 28.6205 | 27.2 | 30.6 | 28.5 | 25.7 | 28.1 | 29.2976 | 29.42 | 27.42 | 28.64 | 21.01 | 30 | 26.2 | 24.11 |  |  |  | 17
| 1 | 3 | 26.87 | 28.77 | 27.4 | 26.9 | 28.1 | 27.6 | 27.7579 | 27.9 | 27.9 | 28 | 25.2 | 27.4 | 29.1174 | 29.08 | 27.68 | 28.87 | 23.82 | 29 | 25.7 | 23.77 |  |  |  | 16
| 1 | 4 | 27.13 | 28.66 | 28.6 | 27.3 | 27 | 26.9 | 26.1972 | 27.4 | 28.6 | 27.9 | 25.7 | 27.9 | 28.7313 | 29.74 | 28.56 | 28.79 | 9.71 | 31 | 24.9 | 24.46 |  |  |  | 27.9047
| 1 | 5 | 27.36 | 29.82 | 28.5 | 27.4 | 26.9 | 27.8 | 26.0111 | 26.7 | 29.3 | 28.5 | 26 | 28.3 | 29.2058 | 29.52 | 28.43 | 28.44 | 23.6 | 30 | 25.3 | 23.88 |  |  |  | 24
| 1 | 6 | 27.28 | 28.35 | 28 | 27.4 | 27.7 | 26.7 | 27.6257 | 28.7 | 29.3 | 28.3 | 26.2 | 26.1 | 28.9074 | 29.53 | 27.95 | 28.95 | 23.63 | 30 | 25.8 | 23.45
| Mean |  | 26.9983 | 28.91 | 28.2 | 27.3167 | 27.5333 | 27.3333 | 27.1781 | 27.55 | 28.9833 | 28.2667 | 25.7833 | 27.75 | 29.034 | 29.4117 | 27.805 | 28.78 | 20.7533 | 29.8333 | 25.6167 | 24.0417
| Median |  | 27 | 28.715 | 28.25 | 27.4 | 27.6 | 27.35 | 27.241 | 27.4 | 28.95 | 28.35 | 25.8 | 28 | 29.031 | 29.47 | 27.815 | 28.83 | 23.175 | 30 | 25.75 | 23.995
| Std Dev. |  | 0.30235 | 0.54743 | 0.49396 | 0.2137 | 0.50067 | 0.50067 | 1.0044 | 0.68337 | 0.97451 | 0.2582 | 0.34303 | 0.91597 | 0.21053 | 0.24384 | 0.65966 | 0.20804 | 5.5098 | 0.75277 | 0.4535 | 0.42883
| Rel.Std.Dev. |  | 1.12% | 1.894% | 1.752% | 0.7823% | 1.818% | 1.832% | 3.696% | 2.48% | 3.362% | 0.9134% | 1.33% | 3.301% | 0.7251% | 0.829% | 2.372% | 0.7229% | 26.55% | 2.523% | 1.77% | 1.784%
| PDM3 |  | -3.248% | 3.603% | 1.058% | -2.107% | -1.331% | -2.047% | -2.604% | -1.271% | 3.866% | 1.297% | -7.602% | -0.5542% | 4.047% | 5.401% | -0.3571% | 3.137% | -25.63% | 6.912% | -8.199% | -13.84%
| Z-Score (Absolute) |  | 0.49 | 0.95 | 0.42 | 0.25 | 0.09 | 0.24 | 0.36 | 0.08 | 1.01 | 0.47 | 1.41 | 0.08 | 1.05 | 1.33 | 0.12 | 0.86 | 5.22 | 1.65 | 1.54 | 2.73
| Analytical results for Gd in OREAS 553b (Certified Value 5.01 ppm) | [24 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 4.46 | 5 | 6 | 8.4584 | 4.83 | 5.4182 | 4.91 | 4.3 | [15 empty] | 1
| 1 | 2 | 4.53 | 4.9 | 6.3 | 8.5032 | 4.77 | 5.1982 | 5.06 | 4.2 | [15 empty] | 18
| 1 | 3 | 4.7 | 5.1 | 6.1 | 8.4736 | 4.83 | 5.4419 | 4.82 | 4.2 | [15 empty] | 16
| 1 | 4 | 4.62 | 4.9 | 6.4 | 8.452 | 4.82 | 5.372 | 4.94 | 4.4 | [15 empty] | 5.006
| 1 | 5 | 4.64 | 5.3 | 6.4 | 8.4368 | 4.94 | 5.3038 | 4.81 | 4 | [15 empty] | 25
| 1 | 6 | 4.69 | 5 | 6 | 8.428 | 4.57 | 5.2997 | 5.18 | 3.6
| Mean |  | 4.6067 | 5.0333 | 6.2 | 8.4587 | 4.7933 | 5.339 | 4.9533 | 4.1167
| Median |  | 4.63 | 5 | 6.2 | 8.4552 | 4.825 | 5.3379 | 4.925 | 4.2
| Std Dev. |  | 0.094163 | 0.15055 | 0.18974 | 0.02709 | 0.12275 | 0.090072 | 0.14362 | 0.28577
| Rel.Std.Dev. |  | 2.044% | 2.991% | 3.06% | 0.3203% | 2.561% | 1.687% | 2.899% | 6.942%
| PDM3 |  | -7.978% | 0.5451% | 23.85% | 68.97% | -4.249% | 6.651% | -1.053% | -17.77%
| Z-Score (Absolute) |  | 0.71 | 0.07 | 2.22 | 6.38 | 0.37 | 0.64 | 0.07 | 1.61
| Analytical results for Ge in OREAS 553b (Indicative Value 0.27 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 18 | 19 | 21 | 25 | 26 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 1.9 | 0.2 | < 0.1 | 1.7 | 0.63 | 0.5 | 0.38473 | 0.17 | 0.21 | 0.16 | 0.18 | 0.23 | 0.412 | 1.4 | 0.1 | 4.4 | 1.93 | < 0.2 | [5 empty] | 1
| 1 | 2 | 1.9 | 0.16 | < 0.1 | 1.7 | 0.62 | 0.4 | 0.34337 | 0.15 | 0.25 | 0.14 | 0.2 | 0.27 | 0.42161 | 1.6 | 0.08 | 4.4 | 1.99 | < 0.2 | [5 empty] | 2
| 1 | 3 | 2 | 0.15 | < 0.1 | 1.7 | 0.65 | 0.5 | 0.35066 | 0.17 | 0.2 | 0.16 | 0.19 | 0.28 | 0.37464 | 1.6 | 0.09 | 4.1 | 1.91 | 0.2 | [5 empty] | 16
| 1 | 4 | 1.9 | 0.3 | < 0.1 | 1.7 | 0.54 | 0.5 | 0.30068 | 0.16 | 0.21 | 0.16 | 0.19 | 0.25 | 0.3821 | 1.6 | 0.03 | 4.6 | 1.88 | 0.2 | [5 empty] | 0.26989
| 1 | 5 | 2 | 0.19 | < 0.1 | 1.7 | 0.59 | 0.4 | 0.2796 | 0.17 | 0.24 | 0.15 | 0.18 | 0.33 | 0.36808 | 1.7 | 0.07 | 4.3 | 1.82 | 0.2 | [5 empty] | 8
| 1 | 6 | 2 | 0.17 | < 0.1 | 1.7 | 0.6 | 0.5 | 0.3297 | 0.15 | 0.2 | 0.15 | 0.18 | 0.26 | 0.36722 | 1.7 | 0.08 | 4.7 | 1.91 | < 0.2
| Mean |  | 1.95 | 0.195 |  | 1.7 | 0.605 | 0.46667 | 0.33146 | 0.16167 | 0.21833 | 0.15333 | 0.18667 | 0.27 | 0.38761 | 1.6 | 0.075 | 4.4167 | 1.9067 | 0.2
| Median |  | 1.95 | 0.18 |  | 1.7 | 0.61 | 0.5 | 0.33653 | 0.165 | 0.21 | 0.155 | 0.185 | 0.265 | 0.37837 | 1.6 | 0.08 | 4.4 | 1.91 | 0.2
| Std Dev. |  | 0.054772 | 0.054681 |  | 0 | 0.038341 | 0.05164 | 0.037384 | 0.0098319 | 0.02137 | 0.008165 | 0.008165 | 0.034059 | 0.023437 | 0.10954 | 0.02429 | 0.2137 | 0.056095 | 3.3993e-17
| Rel.Std.Dev. |  | 2.809% | 28.04% |  | 0% | 6.337% | 11.07% | 11.28% | 6.082% | 9.788% | 5.325% | 4.374% | 12.61% | 6.047% | 6.847% | 32.39% | 4.838% | 2.942% | 1.7e-14%
| PDM3 |  | 622.5% | -27.75% |  | 529.9% | 124.2% | 72.91% | 22.81% | -40.1% | -19.1% | -43.19% | -30.84% | 0.03924% | 43.61% | 492.8% | -72.21% | 1536% | 606.4% | -25.9%
| Z-Score (Absolute) |  | 7.02 | 0.45 | 1.07 | 5.96 | 1.3 | 0.71 | 0.13 | 0.59 | 0.35 | 0.63 | 0.49 | 0.13 | 0.37 | 5.53 | 0.96 | 17.53 | 6.84 | 0.64
| Analytical results for Hf in OREAS 553b (Certified Value 5.46 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 5.48 | 4.69 | 5.8 | 5.5 | 5.48 | 4.9 | 5.54 | 5.3012 | 5.4 | 5.5 | 5.3 | 5.8 | 5.5 | 5.6237 | 5.28 | 5.43 | 5.29 | 4.04 | 4.5 | 5.3 | 5.71 | 5.32 |  | 1
| 1 | 2 | 5.37 | 4.47 | 5.9 | 5.4 | 5.42 | 4.6 | 5.49 | 5.4152 | 5.3 | 5.9 | 5.3 | 5.9 | 5.3 | 5.2836 | 5.28 | 5.26 | 5.35 | 3.74 | 4.5 | 5.2 | 5.83 | 5.36 |  | 22
| 1 | 3 | 5.38 | 4.65 | 5.6 | 5.4 | 5.36 | 4.7 | 5.53 | 5.5011 | 5.6 | 5.5 | 5.2 | 5.7 | 5.4 | 5.506 | 5.26 | 5.57 | 5.17 | 4.51 | 4.5 | 5.1 | 5.77 | 5.41 |  | 16
| 1 | 4 | 5.52 | 4.68 | 5.8 | 5.6 | 5.62 | 4.8 | 5.48 | 5.0823 | 5.4 | 5.5 | 5.2 | 5.7 | 5.3 | 5.4718 | 5.34 | 5.41 | 5.25 | 3.48 | 4.5 | 5.4 | 6.13 | 5.42 |  | 5.4611
| 1 | 5 | 5.51 | 4.77 | 6 | 5.6 | 5.64 | 4.6 | 5.5 | 5.0377 | 5.2 | 5.6 | 5.6 | 5.9 | 5.3 | 5.4295 | 5.31 | 5.06 | 5.36 | 4.56 | 4.5 | 5.2 | 5.45 | 5.38 |  | 26
| 1 | 6 | 5.8 | 4.52 | 5.8 | 5.5 | 5.43 | 4.7 | 5.55 | 5.3423 | 5.6 | 5.7 | 5.3 | 5.9 | 5.2 | 5.3583 | 5.27 | 5.46 | 5.49 | 4.41 | 4.6 | 5 | 5.73 | 5.33
| Mean |  | 5.51 | 4.63 | 5.8167 | 5.5 | 5.4917 | 4.7167 | 5.515 | 5.28 | 5.4167 | 5.6167 | 5.3167 | 5.8167 | 5.3333 | 5.4455 | 5.29 | 5.365 | 5.3183 | 4.1233 | 4.5167 | 5.2 | 5.77 | 5.37
| Median |  | 5.495 | 4.665 | 5.8 | 5.5 | 5.455 | 4.7 | 5.515 | 5.3217 | 5.4 | 5.55 | 5.3 | 5.85 | 5.3 | 5.4506 | 5.28 | 5.42 | 5.32 | 4.225 | 4.5 | 5.2 | 5.75 | 5.37
| Std Dev. |  | 0.15595 | 0.11296 | 0.13292 | 0.089443 | 0.11392 | 0.1169 | 0.02881 | 0.18399 | 0.16021 | 0.16021 | 0.1472 | 0.098319 | 0.10328 | 0.11848 | 0.029665 | 0.17964 | 0.10926 | 0.445 | 0.040825 | 0.14142 | 0.21945 | 0.040988
| Rel.Std.Dev. |  | 2.83% | 2.44% | 2.285% | 1.626% | 2.074% | 2.479% | 0.5224% | 3.485% | 2.958% | 2.852% | 2.769% | 1.69% | 1.936% | 2.176% | 0.5608% | 3.348% | 2.054% | 10.79% | 0.9039% | 2.72% | 3.803% | 0.7633%
| PDM3 |  | 0.8951% | -15.22% | 6.511% | 0.712% | 0.5594% | -13.63% | 0.9867% | -3.317% | -0.8139% | 2.848% | -2.645% | 6.511% | -2.34% | -0.2866% | -3.133% | -1.76% | -2.615% | -24.5% | -17.29% | -4.781% | 5.656% | -1.668%
| Z-Score (Absolute) |  | 0.7 | 3.62 | 2.2 | 0.65 | 0.61 | 3.19 | 0.72 | 0.43 | 0.24 | 1.22 | 0.25 | 2.2 | 0.17 | 0.38 | 0.38 | 0.01 | 0.24 | 6.1 | 4.17 | 0.82 | 1.97 | 0.01
| Analytical results for Hg in OREAS 553b (Indicative Value 0.058 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [20 empty] | 1
| No | No | 15 | 19 | 25 | [20 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [20 empty] | 3
3
| 1 | 1 | < 1 | 0.06 | 0.07 | [20 empty] | 1
| 1 | 2 | < 1 | 0.04 | 0.08 | [20 empty] | 23
| 1 | 3 | < 1 | 0.03 | 0.07 | [20 empty] | 16
| 1 | 4 | < 1 | 0.04 | 0.07 | [20 empty] | 0.058333
| 1 | 5 | < 1 | 0.03 | 0.09 | [20 empty] | 7
| 1 | 6 | < 1 | 0.03 | 0.09
| Mean |  |  | 0.038333 | 0.078333
| Median |  |  | 0.035 | 0.075
| Std Dev. |  |  | 0.01169 | 0.0098319
| Rel.Std.Dev. |  |  | 30.5% | 12.55%
| PDM3 |  |  | -34.29% | 34.29%
| Z-Score (Absolute) |  | 7.11 | 0.67 | 0
| Analytical results for Ho in OREAS 553b (Certified Value 0.96 ppm) | [24 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.89 | 0.7 | 1.1 | 1.143 | 0.96 | 0.96177 | 0.93 | 0.6 | [15 empty] | 1
| 1 | 2 | 0.87 | 0.7 | 1.1 | 1.1344 | 0.86 | 0.93476 | 0.93 | 0.6 | [15 empty] | 5
| 1 | 3 | 0.87 | 0.7 | 1.1 | 1.1551 | 0.92 | 0.99373 | 0.9 | 0.58 | [15 empty] | 16
| 1 | 4 | 0.88 | 0.7 | 1.1 | 1.1255 | 0.91 | 0.97167 | 1 | 0.63 | [15 empty] | 0.96314
| 1 | 5 | 0.89 | 0.8 | 1.1 | 1.1449 | 0.92 | 0.95044 | 0.87 | 0.59 | [15 empty] | 27
| 1 | 6 | 0.89 | 0.7 | 1.1 | 1.1467 | 0.87 | 0.94227 | 0.93 | 0.52
| Mean |  | 0.88167 | 0.71667 | 1.1 | 1.1416 | 0.90667 | 0.95911 | 0.92667 | 0.58667
| Median |  | 0.885 | 0.7 | 1.1 | 1.144 | 0.915 | 0.95611 | 0.93 | 0.595
| Std Dev. |  | 0.0098319 | 0.040825 | 2.4324e-16 | 0.010317 | 0.036697 | 0.021519 | 0.043205 | 0.036697
| Rel.Std.Dev. |  | 1.115% | 5.696% | 2.211e-14% | 0.9038% | 4.047% | 2.244% | 4.662% | 6.255%
| PDM3 |  | -8.46% | -25.59% | 14.21% | 18.53% | -5.864% | -0.419% | -3.787% | -39.09%
| Z-Score (Absolute) |  | 0.61 | NA | NA | 3.92 | 0.17 | 0.74 | 0.17 | 5.75
Results from laboratories 2 and 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 553b (Certified Value 2.14 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.22 | 2.27 | 2.1 | 2.15 | 1.798 | 2.74 | 2.04 | 2.1162 | 2.24 | 2.08 | 2.08 | 2.29 | 2.22 | 2.0698 | 2.15 | 2.111 | 2.24 | 0.362 | 1.92 | 2.05 | 2.25 | 1.787 |  | 1
| 1 | 2 | 2.14 | 2.17 | 2.1 | 2.15 | 1.769 | 2.67 | 2 | 2.213 | 2.22 | 2.29 | 2.04 | 2.28 | 2.14 | 2.0768 | 2.14 | 2.095 | 2.23 | 0.338 | 1.94 | 2.08 | 2.19 | 1.738 |  | 6
| 1 | 3 | 2.22 | 2.31 | 2.1 | 2.1 | 1.817 | 2.87 | 2.03 | 2.1348 | 2.22 | 2.05 | 2.09 | 2.22 | 2.13 | 2.1388 | 2.09 | 2.106 | 2.24 | 0.5 | 2 | 2.13 | 2.22 | 1.769 |  | 16
| 1 | 4 | 2.16 | 2.27 | 2.1 | 2.11 | 1.803 | 2.7 | 1.98 | 2.0684 | 2.2 | 2.15 | 2.02 | 2.27 | 2.12 | 2.0844 | 2.18 | 2.132 | 2.19 | 0.058 | 1.94 | 2.11 | 2.43 | 1.817 |  | 2.1391
| 1 | 5 | 2.24 | 2.45 | 2.2 | 2.15 | 1.808 | 2.67 | 2.06 | 2.0642 | 2.13 | 2.19 | 2.07 | 2.28 | 2.16 | 2.0162 | 2.18 | 2.146 | 2.2 | 0.482 | 1.92 | 2.1 | 2.21 | 1.735 |  | 28
| 1 | 6 | 2.21 | 2.22 | 2.1 | 2.12 | 1.75 | 2.72 | 1.99 | 2.1878 | 2.32 | 2.15 | 2.1 | 2.26 | 2.12 | 2.0132 | 2.16 | 2.144 | 2.19 | 0.478 | 1.96 | 2.08 | 2.21 | 1.742
| Mean |  | 2.1983 | 2.2817 | 2.1167 | 2.13 | 1.7908 | 2.7283 | 2.0167 | 2.1307 | 2.2217 | 2.1517 | 2.0667 | 2.2667 | 2.1483 | 2.0665 | 2.15 | 2.1223 | 2.215 | 0.36967 | 1.9467 | 2.0917 | 2.2517 | 1.7647
| Median |  | 2.215 | 2.27 | 2.1 | 2.135 | 1.8005 | 2.71 | 2.015 | 2.1255 | 2.22 | 2.15 | 2.075 | 2.275 | 2.135 | 2.0733 | 2.155 | 2.1215 | 2.215 | 0.42 | 1.94 | 2.09 | 2.215 | 1.7555
| Std Dev. |  | 0.0392 | 0.095586 | 0.040825 | 0.022804 | 0.025779 | 0.074677 | 0.031411 | 0.060935 | 0.061455 | 0.084951 | 0.030768 | 0.025033 | 0.038166 | 0.046991 | 0.033466 | 0.021285 | 0.02429 | 0.16706 | 0.030111 | 0.027869 | 0.089536 | 0.032745
| Rel.Std.Dev. |  | 1.783% | 4.189% | 1.929% | 1.071% | 1.44% | 2.737% | 1.558% | 2.86% | 2.766% | 3.948% | 1.489% | 1.104% | 1.777% | 2.274% | 1.557% | 1.003% | 1.097% | 45.19% | 1.547% | 1.332% | 3.976% | 1.856%
| PDM3 |  | 2.77% | 6.665% | -1.048% | -0.4248% | -16.28% | 27.55% | -5.723% | -0.3911% | 3.861% | 0.5881% | -3.386% | 5.964% | 0.4323% | -3.393% | 0.5102% | -0.7832% | 3.549% | -82.72% | -8.995% | -2.217% | 5.263% | -17.5%
| Z-Score (Absolute) |  | 0.6 | 1.34 | 0.12 | 0 | 3 | 5.28 | 1 | 0 | 0.81 | 0.19 | 0.56 | 1.2 | 0.16 | 0.56 | 0.17 | 0.07 | 0.75 | 15.56 | 1.62 | 0.34 | 1.07 | 3.23
| Analytical results for K in OREAS 553b (Certified Value 2.55 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.5574 | 1.62 | 2.6 | 2.61 | 2.3093 | 2.83 | 2.57 | 2.47 | 2.5155 | 2.55 | 2.53 | 2.5 | 2.64 | 2.58 | 2.555 | 2.56 | 2.55 | 2.6261 | 2.4907 | 2.7 | 2.45 | 2.56 | 2.66 | 1
| 1 | 2 | 2.5561 | 1.55 | 2.85 | 2.54 | 2.3037 | 2.67 | 2.57 | 2.42 | 2.5129 | 2.54 | 2.55 | 2.46 | 2.66 | 2.54 | 2.5709 | 2.55 | 2.55 | 2.6606 | 2.4728 | 2.7 | 2.43 | 2.59 | 2.66 | #N/A
| 1 | 3 | 2.5744 | 1.68 | 2.26 | 2.49 | 2.305 | 2.76 | 2.58 | 2.53 | 2.5172 | 2.58 | 2.46 | 2.48 | 2.65 | 2.55 | 2.6006 | 2.54 | 2.54 | 2.6178 | 2.4854 | 2.67 | 2.42 | 2.63 | 2.65 | 16
| 1 | 4 | 2.6138 | 2.15 | 2.19 | 2.5 | 2.3069 | 2.5 | 2.55 | 2.42 | 2.511 | 2.55 | 2.47 | 2.5 | 2.67 | 2.55 | 2.6088 | 2.57 | 2.57 | 2.6745 | 2.4312 | 2.67 | 2.44 | 2.54 | 2.68 | 2.5541
| 1 | 5 | 2.636 | 1.28 | 2.23 | 2.6 | 2.3019 | 2.29 | 2.51 | 2.48 | 2.5175 | 2.54 | 2.48 | 2.47 | 2.75 | 2.57 | 2.6025 | 2.57 | 2.56 | 2.6981 | 2.4239 | 2.66 | 2.43 | 2.45 | 2.62 | 29
| 1 | 6 | 2.5644 | 1.43 | 2.23 | 2.59 | 2.3039 | 2.3 | 2.54 | 2.51 | 2.5097 | 2.62 | 2.54 | 2.48 | 2.74 | 2.54 | 2.5503 | 2.57 | 2.54 | 2.7044 | 2.5027 | 2.73 | 2.43 | 2.44 | 2.51
| Mean |  | 2.5837 | 1.6183 | 2.3933 | 2.555 | 2.3051 | 2.5583 | 2.5533 | 2.4717 | 2.514 | 2.5633 | 2.505 | 2.4817 | 2.685 | 2.555 | 2.5813 | 2.56 | 2.5517 | 2.6636 | 2.4678 | 2.6883 | 2.4333 | 2.535 | 2.63
| Median |  | 2.5694 | 1.585 | 2.245 | 2.565 | 2.3045 | 2.585 | 2.56 | 2.475 | 2.5142 | 2.55 | 2.505 | 2.48 | 2.665 | 2.55 | 2.5857 | 2.565 | 2.55 | 2.6676 | 2.4791 | 2.685 | 2.43 | 2.55 | 2.655
| Std Dev. |  | 0.033328 | 0.29701 | 0.26972 | 0.05244 | 0.0026279 | 0.23198 | 0.02582 | 0.04535 | 0.0032537 | 0.031411 | 0.03937 | 0.016021 | 0.047645 | 0.016432 | 0.025865 | 0.012649 | 0.01169 | 0.036018 | 0.032692 | 0.026394 | 0.010328 | 0.076092 | 0.061968
| Rel.Std.Dev. |  | 1.29% | 18.35% | 11.27% | 2.052% | 0.114% | 9.068% | 1.011% | 1.835% | 0.1294% | 1.225% | 1.572% | 0.6456% | 1.774% | 0.6431% | 1.002% | 0.4941% | 0.4581% | 1.352% | 1.325% | 0.9818% | 0.4244% | 3.002% | 2.356%
| PDM3 |  | 1.157% | -36.64% | -6.296% | 0.03392% | -9.75% | 0.1644% | -0.03133% | -3.229% | -1.573% | 0.3602% | -1.924% | -2.837% | 5.124% | 0.03392% | 1.065% | 0.2297% | -0.09659% | 4.285% | -3.381% | 5.254% | -4.73% | -0.7491% | 2.97%
| Z-Score (Absolute) |  | 0.42 | 13.04 | 2.23 | 0.02 | 3.46 | 0.07 | 0 | 1.14 | 0.55 | 0.14 | 0.67 | 1 | 1.84 | 0.02 | 0.39 | 0.09 | 0.02 | 1.54 | 1.19 | 1.88 | 1.67 | 0.26 | 1.07
| Analytical results for La in OREAS 553b (Certified Value 68 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
0
| 1 | 1 | 72.88 | 64 | 76 | 87.56 | 65.47 | 60 | 53.1 | 77.5 | 71.1 | 69.1 | 73.6 | 77.3 | 68.3209 | 75.8 | 71.09 | 77.17 | 55.7 | 47.5 | 62 | 74.8 | 54.8 |  |  | 1
| 1 | 2 | 72.74 | 63.5 | 77.8 | 87.584 | 64.96 | 61 | 51.6 | 76.6 | 77.1 | 84.3 | 73.4 | 75.3 | 68.4965 | 74.4 | 65.85 | 77.09 | 52.1 | 49.9 | 64 | 72.8 | 54.6 |  |  | 7
| 1 | 3 | 72.82 | 62.6 | 74.9 | 87.656 | 65.71 | 62 | 53.5 | 79.5 | 70.9 | 67.9 | 74.5 | 76.7 | 70.6608 | 75.5 | 64.63 | 76.02 | 54 | 49.5 | 64 | 72.6 | 52.3 |  |  | 16
| 1 | 4 | 72.29 | 62.5 | 77.2 | 87.584 | 64.73 | 61 | 51.8 | 76.5 | 62.3 | 67 | 72.4 | 76 | 69.6822 | 75.8 | 67.35 | 78.5 | 21.5 | 54.5 | 64 | 70.8 | 54.4 |  |  | 68.2929
| 1 | 5 | 72.56 | 68.4 | 77.5 | 87.612 | 66.93 | 61 | 53.1 | 74.8 | 72.9 | 70.1 | 71 | 74.5 | 69.3411 | 76.7 | 70.44 | 76.72 | 53.8 | 54.4 | 66 | 71.6 | 49.4 |  |  | 30
| 1 | 6 | 73.84 | 65.1 | 76.4 | 87.608 | 65.86 | 61 | 51.7 | 80.8 | 72.5 | 68.3 | 75.2 | 73.8 | 68.2708 | 76.9 | 66.57 | 82.25 | 52.7 | 58.2 | 64 | 72.2 | 45.1
| Mean |  | 72.855 | 64.35 | 76.6333 | 87.6007 | 65.61 | 61 | 52.4667 | 77.6167 | 71.1333 | 71.1167 | 73.35 | 75.6 | 69.1287 | 75.85 | 67.655 | 77.9583 | 48.3 | 52.3333 | 64 | 72.4667 | 51.7667
| Median |  | 72.78 | 63.75 | 76.8 | 87.596 | 65.59 | 61 | 52.45 | 77.05 | 71.8 | 68.7 | 73.5 | 75.65 | 68.9188 | 75.8 | 66.96 | 77.13 | 53.25 | 52.15 | 64 | 72.4 | 53.35
| Std Dev. |  | 0.52755 | 2.2043 | 1.0857 | 0.033025 | 0.77864 | 0.63246 | 0.85479 | 2.1867 | 4.8718 | 6.5441 | 1.499 | 1.3266 | 0.94732 | 0.90056 | 2.5782 | 2.2529 | 13.1874 | 4.0173 | 1.2649 | 1.3545 | 3.8599
| Rel.Std.Dev. |  | 0.7241% | 3.426% | 1.417% | 0.0377% | 1.187% | 1.037% | 1.629% | 2.817% | 6.849% | 9.202% | 2.044% | 1.755% | 1.37% | 1.187% | 3.811% | 2.89% | 27.3% | 7.676% | 1.976% | 1.869% | 7.456%
| PDM3 |  | 6.68% | -5.774% | 12.21% | 28.27% | -3.929% | -10.68% | -23.17% | 13.65% | 4.159% | 4.135% | 7.405% | 10.7% | 1.224% | 11.07% | -0.9341% | 14.15% | -29.28% | -23.37% | -6.286% | 6.112% | -24.2%
| Z-Score (Absolute) |  | 0.21 | 0.83 | 0.67 | 2.01 | 0.67 | 1.24 | 2.28 | 0.79 | 0 | 0 | 0.27 | 0.55 | 0.24 | 0.58 | 0.42 | 0.84 | 2.79 | 2.3 | 0.87 | 0.17 | 2.37
| Analytical results for Li in OREAS 553b (Certified Value 63 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 61.1 | 96.6 | 61 | 65.1 | 65.43 | 81 | 64 | 57 | 63.5475 | 64.3 | 62 | 65 | 64.3 | 65.4 | 61.1488 | 61.9 | 61.2 | 63.7 | 61 | 63 | 59 | 60.6 | 63 | 1
| 1 | 2 | 60.7 | 95.4 | 60 | 64.1 | 66.01 | 77 | 64 | 59 | 67.0734 | 64.2 | 62.7 | 64 | 63 | 64.1 | 60.861 | 61.6 | 61.2 | 63.6 | 61 | 61.4 | 63.9 | 61.3 | 63 | #N/A
| 1 | 3 | 61.6 | 96.6 | 61 | 64.1 | 65.92 | 80 | 64 | 56 | 67.3234 | 64.8 | 60.6 | 64.5 | 62.8 | 64.7 | 62.6367 | 61.3 | 60.9 | 63.9 | 60 | 61.6 | 61.1 | 59.4 | 68 | 16
| 1 | 4 | 61.3 | 97.3 | 62 | 64.1 | 66.19 | 71 | 64 | 59 | 63.1251 | 64.1 | 61.3 | 65 | 63.8 | 64.4 | 62.8007 | 63.5 | 61.8 | 63.8 | 59 | 60.6 | 61.6 | 59.9 | 68 | 62.6357
| 1 | 5 | 62.6 | 100.8 | 60 | 65.1 | 66.48 | 64 | 63 | 55 | 62.5167 | 64.2 | 60.8 | 64.4 | 65 | 64.7 | 62.0269 | 61.3 | 61.8 | 64.2 | 59 | 59.8 | 59.6 | 58.4 | 66 | 31
| 1 | 6 | 60.9 | 95.1 | 61 | 64.1 | 66.2 | 64 | 64 | 58 | 65.0946 | 65.9 | 62.7 | 64.2 | 66.3 | 64.3 | 61.3119 | 61.7 | 61.2 | 64.5 | 62 | 61.5 | 60.7 | 58.8 | 64
| Mean |  | 61.3667 | 96.9667 | 60.8333 | 64.4333 | 66.0383 | 72.8333 | 63.8333 | 57.3333 | 64.7801 | 64.5833 | 61.6833 | 64.5167 | 64.2 | 64.6 | 61.7976 | 61.8833 | 61.35 | 63.95 | 60.3333 | 61.3167 | 60.9833 | 59.7333 | 65.3333
| Median |  | 61.2 | 96.6 | 61 | 64.1 | 66.1 | 74 | 64 | 57.5 | 64.3211 | 64.25 | 61.65 | 64.45 | 64.05 | 64.55 | 61.6694 | 61.65 | 61.2 | 63.85 | 60.5 | 61.45 | 60.9 | 59.65 | 65
| Std Dev. |  | 0.6802 | 2.05 | 0.75277 | 0.5164 | 0.35471 | 7.679 | 0.40825 | 1.633 | 2.0595 | 0.69113 | 0.92394 | 0.41191 | 1.313 | 0.45607 | 0.81205 | 0.82563 | 0.36742 | 0.33912 | 1.2111 | 1.0741 | 1.7221 | 1.0948 | 2.3381
| Rel.Std.Dev. |  | 1.108% | 2.114% | 1.237% | 0.8014% | 0.5371% | 10.54% | 0.6396% | 2.848% | 3.179% | 1.07% | 1.498% | 0.6384% | 2.045% | 0.706% | 1.314% | 1.334% | 0.5989% | 0.5303% | 2.007% | 1.752% | 2.824% | 1.833% | 3.579%
| PDM3 |  | -2.026% | 54.81% | -2.878% | 2.87% | 5.432% | 16.28% | 1.912% | -8.465% | 3.424% | 3.109% | -1.52% | 3.003% | 2.497% | 3.136% | -1.338% | -1.201% | -2.053% | 2.098% | -3.676% | -2.106% | -2.638% | -4.634% | 4.307%
| Z-Score (Absolute) |  | 0.77 | 10.39 | 0.94 | 0.19 | 0.69 | 2.82 | 0 | 2.04 | 0.3 | 0.24 | 0.67 | 0.21 | 0.11 | 0.24 | 0.64 | 0.61 | 0.78 | 0.04 | 1.1 | 0.79 | 0.89 | 1.29 | 0.47
| Analytical results for Lu in OREAS 553b (Certified Value 0.42 ppm) | [24 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 15 | 17 | 25 | 26 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
2
| 1 | 1 | 0.43 | 0.4 | 0.4 | 0.58984 | 0.51 | 0.39 | 0.45735 | 0.41 | 0.44 | 0.32 | [13 empty] | 1
| 1 | 2 | 0.42 | 0.4 | 0.5 | 0.5352 | 0.47 | 0.36 | 0.4413 | 0.43 | 0.43 | 0.32 | [13 empty] | 8
| 1 | 3 | 0.41 | 0.4 | 0.4 | 0.55848 | 0.49 | 0.38 | 0.45304 | 0.42 | 0.43 | 0.32 | [13 empty] | 16
| 1 | 4 | 0.42 | 0.4 | 0.5 | 0.59288 | 0.49 | 0.37 | 0.46732 | 0.45 | 0.48 | 0.32 | [13 empty] | 0.41808
| 1 | 5 | 0.43 | 0.4 | 0.5 | 0.58424 | 0.49 | 0.38 | 0.45063 | 0.45 | 0.41 | 0.31 | [13 empty] | 32
| 1 | 6 | 0.43 | 0.4 | 0.5 | 0.57632 | 0.47 | 0.36 | 0.4518 | 0.44 | 0.44 | 0.28
| Mean |  | 0.42333 | 0.4 | 0.46667 | 0.57283 | 0.48667 | 0.37333 | 0.45357 | 0.43333 | 0.43833 | 0.31167
| Median |  | 0.425 | 0.4 | 0.5 | 0.58028 | 0.49 | 0.375 | 0.45242 | 0.435 | 0.435 | 0.32
| Std Dev. |  | 0.008165 | 6.0809e-17 | 0.05164 | 0.022161 | 0.015055 | 0.012111 | 0.0085565 | 0.01633 | 0.023166 | 0.016021
| Rel.Std.Dev. |  | 1.929% | 1.52e-14% | 11.07% | 3.869% | 3.094% | 3.244% | 1.886% | 3.768% | 5.285% | 5.14%
| PDM3 |  | 1.256% | -4.325% | 11.62% | 37.01% | 16.4% | -10.7% | 8.489% | 3.648% | 4.844% | -25.45%
| Z-Score (Absolute) |  | 0.25 | NA | NA | 2.69 | 1 | 1.23 | 0.35 | 0.05 | 0.05 | 2.44
Results from laboratories 2 and 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 553b (Certified Value 3.58 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 3.5575 | 3.69 | 3.82 | 3.66 | 2.9163 | 4.089 | 3.56 | 3.35 | 3.6278 | 3.56 | 3.54 | 3.57 | 3.64 | 3.64 | 3.7049 | 3.86 | 3.547 | 3.5894 | 3.7022 | 3.44 | 3.25 | 3.26 | 3.668 | 1
| 1 | 2 | 3.5352 | 3.63 | 3.87 | 3.54 | 2.9153 | 3.899 | 3.57 | 3.43 | 3.6004 | 3.55 | 3.56 | 3.5 | 3.6 | 3.55 | 3.6491 | 3.83 | 3.546 | 3.5923 | 3.6605 | 3.46 | 3.3 | 3.34 | 3.659 | #N/A
| 1 | 3 | 3.5557 | 3.64 | 3.74 | 3.49 | 2.9161 | 3.98 | 3.58 | 3.33 | 3.6522 | 3.61 | 3.41 | 3.54 | 3.58 | 3.58 | 3.6989 | 3.92 | 3.536 | 3.6249 | 3.6657 | 3.45 | 3.33 | 3.37 | 3.653 | 16
| 1 | 4 | 3.6309 | 3.6 | 3.83 | 3.53 | 2.9242 | 3.634 | 3.54 | 3.41 | 3.626 | 3.56 | 3.45 | 3.58 | 3.62 | 3.55 | 3.6936 | 3.83 | 3.577 | 3.6365 | 3.5657 | 3.45 | 3.38 | 3.26 | 3.686 | 3.5765
| 1 | 5 | 3.6572 | 3.85 | 3.82 | 3.65 | 2.9211 | 3.257 | 3.49 | 3.36 | 3.6031 | 3.55 | 3.46 | 3.54 | 3.69 | 3.58 | 3.7087 | 3.87 | 3.579 | 3.5944 | 3.5965 | 3.45 | 3.32 | 3.15 | 3.573 | 33
| 1 | 6 | 3.5784 | 3.65 | 3.7 | 3.56 | 2.9194 | 3.315 | 3.53 | 3.39 | 3.6348 | 3.65 | 3.57 | 3.54 | 3.75 | 3.58 | 3.6445 | 3.89 | 3.55 | 3.6434 | 3.7359 | 3.56 | 3.26 | 3.28 | 3.38
| Mean |  | 3.5858 | 3.6767 | 3.7967 | 3.5717 | 2.9187 | 3.6957 | 3.545 | 3.3783 | 3.624 | 3.58 | 3.4983 | 3.545 | 3.6467 | 3.58 | 3.6833 | 3.8667 | 3.5558 | 3.6135 | 3.6544 | 3.4683 | 3.3067 | 3.2767 | 3.6032
| Median |  | 3.568 | 3.645 | 3.82 | 3.55 | 2.9179 | 3.7665 | 3.55 | 3.375 | 3.6269 | 3.56 | 3.5 | 3.54 | 3.63 | 3.58 | 3.6963 | 3.865 | 3.5485 | 3.6097 | 3.6631 | 3.45 | 3.31 | 3.27 | 3.656
| Std Dev. |  | 0.047862 | 0.089815 | 0.063456 | 0.068532 | 0.0034806 | 0.35158 | 0.032711 | 0.038166 | 0.019616 | 0.040988 | 0.066758 | 0.028107 | 0.06314 | 0.032863 | 0.028775 | 0.035024 | 0.017815 | 0.024282 | 0.063737 | 0.04535 | 0.048028 | 0.076594 | 0.11609
| Rel.Std.Dev. |  | 1.335% | 2.443% | 1.671% | 1.919% | 0.1193% | 9.513% | 0.9227% | 1.13% | 0.5413% | 1.145% | 1.908% | 0.7929% | 1.731% | 0.918% | 0.7812% | 0.9058% | 0.501% | 0.672% | 1.744% | 1.308% | 1.452% | 2.338% | 3.222%
| PDM3 |  | 0.26% | 2.8% | 6.155% | -0.1357% | -18.39% | 3.331% | -0.8813% | -5.541% | 1.329% | 0.09733% | -2.186% | -0.8813% | 1.961% | 0.09733% | 2.985% | 8.113% | -0.5784% | 1.034% | 2.178% | -3.025% | -7.545% | -8.384% | 0.7451%
| Z-Score (Absolute) |  | 0.05 | 0.88 | 1.96 | 0.08 | 5.99 | 1.05 | 0.32 | 1.83 | 0.4 | 0 | 0.74 | 0.32 | 0.6 | 0 | 0.94 | 2.6 | 0.22 | 0.3 | 0.67 | 1.01 | 2.48 | 2.75 | 0.21
| Analytical results for Mn in OREAS 553b (Certified Value 0.013 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0131 | 0.0136 | 0.0123 | 0.0135 | 0.013525 | 0.0165 | 0.0139 | 0.0121 | 0.01334 | 0.0133 | 0.0133 | 0.0133 | 0.0135 | 0.0134 | 0.014102 | 0.0137 | 0.0136 | 0.0146 | 0.0126 | 0.0131 | 0.0138 | 0.0068 | 0.0137 | 1
| 1 | 2 | 0.0131 | 0.0135 | 0.0119 | 0.0134 | 0.013374 | 0.0159 | 0.0139 | 0.0128 | 0.013481 | 0.0131 | 0.0134 | 0.013 | 0.0133 | 0.0131 | 0.013913 | 0.0137 | 0.0138 | 0.0145 | 0.0124 | 0.0134 | 0.0138 | 0.0073 | 0.0136 | #N/A
| 1 | 3 | 0.0131 | 0.0138 | 0.0115 | 0.0134 | 0.013385 | 0.0162 | 0.0144 | 0.0119 | 0.01362 | 0.0134 | 0.0129 | 0.0133 | 0.0133 | 0.0131 | 0.014186 | 0.0137 | 0.0136 | 0.0145 | 0.0124 | 0.0133 | 0.0137 | 0.0074 | 0.0135 | 16
| 1 | 4 | 0.0133 | 0.0135 | 0.0124 | 0.0132 | 0.013364 | 0.0148 | 0.0135 | 0.0126 | 0.013378 | 0.0131 | 0.013 | 0.0134 | 0.0133 | 0.0131 | 0.013933 | 0.0138 | 0.0138 | 0.0145 | 0.0122 | 0.0133 | 0.0137 | 0.0067 | 0.0137 | 0.013474
| 1 | 5 | 0.0136 | 0.0146 | 0.0122 | 0.0132 | 0.013486 | 0.0134 | 0.0131 | 0.0119 | 0.01346 | 0.0132 | 0.013 | 0.0133 | 0.0137 | 0.0136 | 0.013958 | 0.0134 | 0.0138 | 0.0145 | 0.0122 | 0.0131 | 0.0136 | 0.0072 | 0.0131 | 34
| 1 | 6 | 0.0131 | 0.0137 | 0.0121 | 0.0132 | 0.01331 | 0.0135 | 0.0142 | 0.0123 | 0.013651 | 0.0135 | 0.0135 | 0.0133 | 0.0139 | 0.0132 | 0.013773 | 0.0136 | 0.0136 | 0.0148 | 0.0127 | 0.0134 | 0.0136 | 0.0075 | 0.0123
| Mean |  | 0.013217 | 0.013783 | 0.012067 | 0.013317 | 0.013407 | 0.01505 | 0.013833 | 0.012267 | 0.013488 | 0.013267 | 0.013183 | 0.013267 | 0.0135 | 0.01325 | 0.013978 | 0.01365 | 0.0137 | 0.014567 | 0.012417 | 0.013267 | 0.0137 | 0.00715 | 0.013317
| Median |  | 0.0131 | 0.01365 | 0.01215 | 0.0133 | 0.013379 | 0.01535 | 0.0139 | 0.0122 | 0.013471 | 0.01325 | 0.01315 | 0.0133 | 0.0134 | 0.01315 | 0.013945 | 0.0137 | 0.0137 | 0.0145 | 0.0124 | 0.0133 | 0.0137 | 0.00725 | 0.01355
| Std Dev. |  | 0.00020412 | 0.00041673 | 0.0003266 | 0.00013292 | 8.1235e-05 | 0.0013664 | 0.00047188 | 0.00037238 | 0.00012564 | 0.0001633 | 0.00024833 | 0.00013663 | 0.00025298 | 0.00020736 | 0.0001465 | 0.00013784 | 0.00010954 | 0.00012111 | 0.00020412 | 0.00013663 | 8.9443e-05 | 0.00032711 | 0.00054559
| Rel.Std.Dev. |  | 1.544% | 3.023% | 2.707% | 0.9981% | 0.6059% | 9.079% | 3.411% | 3.036% | 0.9315% | 1.231% | 1.884% | 1.03% | 1.874% | 1.565% | 1.048% | 1.01% | 0.7996% | 0.8314% | 1.644% | 1.03% | 0.6529% | 4.575% | 4.097%
| PDM3 |  | -1.906% | 2.299% | -10.44% | -1.164% | -0.4913% | 11.7% | 2.67% | -8.957% | 0.1102% | -1.535% | -2.154% | -1.535% | 0.1965% | -1.659% | 3.741% | 1.31% | 1.681% | 8.113% | -7.844% | -1.535% | 1.681% | -46.93% | -1.164%
| Z-Score (Absolute) |  | 0.2 | 0.94 | 2.53 | 0 | 0.18 | 3.51 | 1.05 | 2.12 | 0.35 | 0.1 | 0.27 | 0.1 | 0.37 | 0.13 | 1.34 | 0.67 | 0.78 | 2.53 | 1.82 | 0.1 | 0.78 | 12.47 | 0
| Analytical results for Mo in OREAS 553b (Certified Value 15.1 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 15.5 | 14.6 | 16 | 15.1 | 14.2 | 15.4 | 12.4554 | 15.2 | 14.1 | 15.3 | 15.2 | 14.35 | 15.1415 | 14.4 | 15.67 | 15.6 | 15.7 | 13.6 | 15.1 | 14.6 | 14.2 |  |  | 1
| 1 | 2 | 15.2 | 13.61 | 17 | 15.3 | 15 | 15.11 | 13.2089 | 15.05 | 15.25 | 15.1 | 14.5 | 13.9 | 15.4193 | 14.13 | 16.01 | 15.8 | 15.5 | 13.4 | 15.2 | 14.2 | 14.2 |  |  | 9
| 1 | 3 | 15.6 | 14.54 | 16 | 15.2 | 15.7 | 15.33 | 13.1826 | 15.4 | 14.25 | 14.95 | 14.85 | 14 | 15.527 | 14.29 | 15.62 | 15.7 | 15.4 | 13.6 | 15.2 | 14.4 | 14.4 |  |  | 16
| 1 | 4 | 15.4 | 14.65 | 16 | 15.2 | 14.9 | 15.14 | 12.0187 | 15.05 | 14.95 | 15.1 | 14.9 | 13.85 | 15.2589 | 14.31 | 16.28 | 15.7 | 15.2 | 13.3 | 15.7 | 15.6 | 14.5 |  |  | 15.0823
| 1 | 5 | 15.4 | 15.44 | 16 | 15.3 | 15.3 | 15.43 | 11.7495 | 14.7 | 14.95 | 15.2 | 15.4 | 13.75 | 15.0194 | 14.19 | 16.02 | 15.7 | 15.3 | 13.3 | 15.3 | 13.6 | 14.3 |  |  | 35
| 1 | 6 | 15.2 | 14.19 | 17 | 15.3 | 15.4 | 15.24 | 12.5307 | 15.95 | 15.15 | 15.4 | 15.05 | 13.8 | 15.242 | 14.66 | 15.71 | 15.6 | 15.9 | 13.6 | 15.4 | 13.8 | 14.2
| Mean |  | 15.3833 | 14.505 | 16.3333 | 15.2333 | 15.0833 | 15.275 | 12.5243 | 15.225 | 14.775 | 15.175 | 14.9833 | 13.9417 | 15.268 | 14.33 | 15.885 | 15.6833 | 15.5 | 13.4667 | 15.3167 | 14.3667 | 14.3
| Median |  | 15.4 | 14.57 | 16 | 15.25 | 15.15 | 15.285 | 12.4931 | 15.125 | 14.95 | 15.15 | 14.975 | 13.875 | 15.2504 | 14.3 | 15.86 | 15.7 | 15.45 | 13.5 | 15.25 | 14.3 | 14.25
| Std Dev. |  | 0.16021 | 0.60096 | 0.5164 | 0.08165 | 0.51929 | 0.13368 | 0.59378 | 0.42279 | 0.4814 | 0.16047 | 0.31091 | 0.21775 | 0.18362 | 0.1873 | 0.2596 | 0.075277 | 0.26077 | 0.15055 | 0.2137 | 0.70899 | 0.12649
| Rel.Std.Dev. |  | 1.041% | 4.143% | 3.162% | 0.536% | 3.443% | 0.8751% | 4.741% | 2.777% | 3.258% | 1.057% | 2.075% | 1.562% | 1.203% | 1.307% | 1.634% | 0.48% | 1.682% | 1.118% | 1.395% | 4.935% | 0.8846%
| PDM3 |  | 1.996% | -3.827% | 8.295% | 1.002% | 0.007088% | 1.278% | -16.96% | 0.9464% | -2.037% | 0.6149% | -0.6559% | -7.563% | 1.232% | -4.988% | 5.322% | 3.985% | 2.77% | -10.71% | 1.554% | -4.745% | -5.187%
| Z-Score (Absolute) |  | 0.35 | 1.13 | 1.95 | 0.1 | 0.15 | 0.17 | 4.47 | 0.08 | 0.67 | 0 | 0.32 | 2.08 | 0.16 | 1.42 | 1.2 | 0.86 | 0.55 | 2.88 | 0.24 | 1.36 | 1.48
| Analytical results for Na in OREAS 553b (Certified Value 0.045 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0437 | 0.048 | 0.05 | 0.034 | 0.044 | 0.04 | 0.32 | 0.04941 | 0.05 | 0.04 | 0.05 | 0.05 | 0.06 | 0.046425 | 0.048 | 0.045 | 0.0449 | 0.042 | 0.04 | 0.04 | 0.035 | 0.052 |  | 1
| 1 | 2 | 0.0438 | 0.045 | 0.05 | 0.035 | 0.033 | 0.04 | 0.32 | 0.047767 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.046736 | 0.05 | 0.046 | 0.0452 | 0.0414 | 0.04 | 0.05 | 0.0361 | 0.052 |  | 20
| 1 | 3 | 0.0437 | 0.047 | 0.05 | 0.035 | 0.04 | 0.04 | 0.3 | 0.046593 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.046176 | 0.049 | 0.046 | 0.0455 | 0.0417 | 0.041 | 0.05 | 0.0344 | 0.049 |  | 16
| 1 | 4 | 0.0446 | 0.049 | 0.05 | 0.035 | 0.028 | 0.04 | 0.3 | 0.043926 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.046896 | 0.049 | 0.046 | 0.0452 | 0.0406 | 0.041 | 0.06 | 0.0384 | 0.05 |  | 0.044918
| 1 | 5 | 0.0447 | 0.047 | 0.05 | 0.034 | 0.024 | 0.05 | 0.31 | 0.043364 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.0461 | 0.049 | 0.045 | 0.0444 | 0.0407 | 0.04 | 0.06 | 0.0335 | 0.05 |  | 36
| 1 | 6 | 0.0437 | 0.045 | 0.05 | 0.034 | 0.025 | 0.04 | 0.3 | 0.046657 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.045286 | 0.049 | 0.045 | 0.0454 | 0.0423 | 0.041 | 0.05 | 0.0349 | 0.05
| Mean |  | 0.044033 | 0.046833 | 0.05 | 0.0345 | 0.032333 | 0.041667 | 0.30833 | 0.046286 | 0.046667 | 0.04 | 0.05 | 0.05 | 0.051667 | 0.04627 | 0.049 | 0.0455 | 0.0451 | 0.04145 | 0.0405 | 0.051667 | 0.035383 | 0.0505
| Median |  | 0.04375 | 0.047 | 0.05 | 0.0345 | 0.0305 | 0.04 | 0.305 | 0.046625 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.046301 | 0.049 | 0.0455 | 0.0452 | 0.04155 | 0.0405 | 0.05 | 0.03495 | 0.05
| Std Dev. |  | 0.00048028 | 0.0016021 | 7.6012e-18 | 0.00054772 | 0.0082138 | 0.0040825 | 0.0098319 | 0.0022932 | 0.005164 | 0 | 7.6012e-18 | 7.6012e-18 | 0.0040825 | 0.00057274 | 0.00063246 | 0.00054772 | 0.0004 | 0.0006892 | 0.00054772 | 0.0075277 | 0.0017034 | 0.0012247
| Rel.Std.Dev. |  | 1.091% | 3.421% | 1.52e-14% | 1.588% | 25.4% | 9.798% | 3.189% | 4.954% | 11.07% | 0% | 1.52e-14% | 1.52e-14% | 7.902% | 1.238% | 1.291% | 1.204% | 0.8869% | 1.663% | 1.352% | 14.57% | 4.814% | 2.425%
| PDM3 |  | -1.969% | 4.265% | 11.31% | -23.19% | -28.02% | -7.238% | 586.4% | 3.047% | 3.894% | -10.95% | 11.31% | 11.31% | 15.03% | 3.01% | 9.088% | 1.296% | 0.4057% | -7.72% | -9.835% | 15.03% | -21.23% | 12.43%
| Z-Score (Absolute) |  | 0.38 | 0.09 | 0.63 | 2 | 2.37 | 0.78 | 44.49 | 0 | 0.07 | 1.07 | 0.63 | 0.63 | 0.91 | 0 | 0.46 | 0.13 | 0.2 | 0.82 | 0.98 | 0.91 | 1.85 | 0.72
| Analytical results for Nb in OREAS 553b (Certified Value 17.4 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
1
| 1 | 1 | 18.97 | 20.43 | 28.5 | 9 | 8.3 | 8.5 | 12.4748 | 18.7 | 19.5 | 20.1 | 18.2 | 21.8 | 19.3494 | 20.6 | 13.6 | 15.62 | NA | 10.8 | 15 | 16.3 | 16.1 |  |  | 1
| 1 | 2 | 19.04 | 18.7 | 29 | 10.1 | 8.4 | 7.5 | 12.0897 | 18.6 | 20.9 | 20.1 | 17.6 | 21.1 | 19.121 | 20.1 | 14.4 | 16.48 | NA | 11 | 14.3 | 18.7 | 15.4 |  |  | 26
| 1 | 3 | 19.94 | 18.67 | 27.4 | 8.7 | 9.1 | 8.6 | 12.4364 | 18.8 | 19.7 | 19.6 | 17.3 | 21 | 19.4884 | 20.4 | 14.4 | 16.24 | NA | 11.2 | 15.8 | 17.5 | 16.8 |  |  | 16
| 1 | 4 | 20.28 | 19.44 | 28.3 | 8.4 | 8.5 | 7.7 | 10.8242 | 18.6 | 19.2 | 19.8 | 18 | 21.7 | 19.1551 | 20.5 | 14 | 16.04 | NA | 11.5 | 15.2 | 16.6 | 16.3 |  |  | 17.3826
| 1 | 5 | 18.75 | 19.49 | 28.6 | 10 | 8.6 | 8.6 | 11.1267 | 18.3 | 19.9 | 20.6 | 17.6 | 21.1 | 19.2231 | 21.1 | 14.7 | 16.65 | NA | 11.1 | 15.1 | 17.9 | 17.3 |  |  | 37
| 1 | 6 | 19.03 | 18.93 | 27.9 | 8.5 | 9.2 | 7.5 | 11.7931 | 19.7 | 19.7 | 20.1 | 17.1 | 21 | 18.8812 | 21 | 13.5 | 14.67 | NA | 11.1 | 15 | 17.5 | 18.2
| Mean |  | 19.335 | 19.2767 | 28.2833 | 9.1167 | 8.6833 | 8.0667 | 11.7908 | 18.7833 | 19.8167 | 20.05 | 17.6333 | 21.2833 | 19.203 | 20.6167 | 14.1 | 15.95 |  | 11.1167 | 15.0667 | 17.4167 | 16.6833
| Median |  | 19.035 | 19.185 | 28.4 | 8.85 | 8.55 | 8.1 | 11.9414 | 18.65 | 19.7 | 20.1 | 17.6 | 21.1 | 19.1891 | 20.55 | 14.2 | 16.14 |  | 11.1 | 15.05 | 17.5 | 16.55
| Std Dev. |  | 0.6188 | 0.6668 | 0.56362 | 0.75211 | 0.37639 | 0.55377 | 0.6855 | 0.47924 | 0.58109 | 0.33912 | 0.41312 | 0.3656 | 0.20779 | 0.37639 | 0.48166 | 0.72255 |  | 0.23166 | 0.48028 | 0.87274 | 0.98268
| Rel.Std.Dev. |  | 3.2% | 3.459% | 1.993% | 8.25% | 4.335% | 6.865% | 5.814% | 2.551% | 2.932% | 1.691% | 2.343% | 1.718% | 1.082% | 1.826% | 3.416% | 4.53% |  | 2.084% | 3.188% | 5.011% | 5.89%
| PDM3 |  | 11.23% | 10.9% | 62.71% | -47.55% | -50.05% | -53.59% | -32.17% | 8.058% | 14% | 15.34% | 1.442% | 22.44% | 10.47% | 18.6% | -18.88% | -8.242% |  | -36.05% | -13.32% | 0.1958% | -4.023%
| Z-Score (Absolute) |  | 0.49 | 0.47 | 2.91 | 2.28 | 2.39 | 2.56 | 1.55 | 0.34 | 0.62 | 0.68 | 0.03 | 1.02 | 0.45 | 0.84 | 0.93 | 0.43 | NA | 1.73 | 0.67 | 0.03 | 0.23
| Analytical results for Nd in OREAS 553b (Certified Value 63 ppm) | [24 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 0
0
| 1 | 1 | 58.78 | 52 | 64.6 | 75.1096 | 61.3 | 63.8646 | 64.1 | 42.3 | [15 empty] | 1
| 1 | 2 | 59.92 | 50.8 | 65.3 | 75.128 | 59.8 | 61.7347 | 62.4 | 43.3 | [15 empty] | 12
| 1 | 3 | 60.1 | 52.2 | 64.4 | 75.1888 | 59.7 | 65.6009 | 61.3 | 41.1 | [15 empty] | 16
| 1 | 4 | 60.41 | 50.2 | 65.8 | 75.14 | 59.3 | 64.4487 | 69.4 | 43.5 | [15 empty] | 62.7385
| 1 | 5 | 61.34 | 53.7 | 65.6 | 75.184 | 60.3 | 63.4098 | 60.7 | 39.9 | [15 empty] | 38
| 1 | 6 | 61.45 | 50.3 | 64.6 | 75.1256 | 59.3 | 64.0831 | 61.9 | 34.1
| Mean |  | 60.3333 | 51.5333 | 65.05 | 75.146 | 59.95 | 63.857 | 63.3 | 40.7
| Median |  | 60.255 | 51.4 | 64.95 | 75.134 | 59.75 | 63.9739 | 62.15 | 41.7
| Std Dev. |  | 0.99039 | 1.356 | 0.59245 | 0.032793 | 0.75829 | 1.2765 | 3.2056 | 3.5077
| Rel.Std.Dev. |  | 1.642% | 2.631% | 0.9108% | 0.04364% | 1.265% | 1.999% | 5.064% | 8.618%
| PDM3 |  | -3.834% | -17.86% | 3.684% | 19.78% | -4.445% | 1.783% | 0.895% | -35.13%
| Z-Score (Absolute) |  | 0.38 | 2.63 | 0.83 | 3.41 | 0.48 | 0.52 | 0.38 | 5.4
| Analytical results for Ni in OREAS 553b (Certified Value 124 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 125.7 | 113.1 | 127 | 125 | 121.15 | 122 | 126.7 | 122.9931 | 121 | 123.5 | 128.5 | 131 | 128 | 118.3404 | 126 | 125 | 130.5 | 127.4 | 122.1 | 113.8 | 121 | 127 |  | 1
| 1 | 2 | 122.4 | 111.6 | 124 | 125 | 122.22 | 122 | 123.1 | 120.8796 | 121 | 129 | 127 | 130 | 124 | 120.5029 | 124 | 124 | 130.3 | 126.2 | 121.7 | 112.9 | 120 | 123 |  | 29
| 1 | 3 | 123.3 | 111.9 | 123 | 125 | 122.09 | 122 | 126.5 | 123.0879 | 121.5 | 121 | 127 | 128 | 126.5 | 121.0291 | 124 | 125 | 129.6 | 125.4 | 120.5 | 113.9 | 121 | 124 |  | 16
| 1 | 4 | 126.3 | 112.7 | 127 | 125 | 118.84 | 122 | 124.3 | 121.7805 | 119.5 | 124.5 | 129 | 131.5 | 124 | 117.6899 | 127 | 126 | 129.7 | 122.8 | 120.5 | 113.9 | 135 | 127 |  | 124.4568
| 1 | 5 | 125.6 | 117.4 | 128 | 125 | 119.77 | 123 | 127.3 | 120.6735 | 119.5 | 124.5 | 127 | 133.5 | 127 | 118.9778 | 126 | 126 | 130.9 | 124 | 119.8 | 114.1 | 117 | 125 |  | 39
| 1 | 6 | 125.9 | 111.7 | 128 | 126 | 119.28 | 122 | 125.8 | 122.1945 | 124 | 127 | 127.5 | 133 | 122.5 | 118.3178 | 127 | 125 | 131.1 | 128.3 | 124.5 | 115.7 | 120 | 125
| Mean |  | 124.8667 | 113.0667 | 126.1667 | 125.1667 | 120.5583 | 122.1667 | 125.6167 | 121.9349 | 121.0833 | 124.9167 | 127.6667 | 131.1667 | 125.3333 | 119.143 | 125.6667 | 125.1667 | 130.35 | 125.6833 | 121.5167 | 114.05 | 122.3333 | 125.1667
| Median |  | 125.65 | 112.3 | 127 | 125 | 120.46 | 122 | 126.15 | 121.9875 | 121 | 124.5 | 127.25 | 131.25 | 125.25 | 118.6591 | 126 | 125 | 130.4 | 125.8 | 121.1 | 113.9 | 120.5 | 125
| Std Dev. |  | 1.6058 | 2.2042 | 2.137 | 0.40825 | 1.4605 | 0.40825 | 1.6055 | 1.0242 | 1.6558 | 2.7824 | 0.8756 | 2.0166 | 2.137 | 1.332 | 1.3663 | 0.75277 | 0.61237 | 2.0634 | 1.6905 | 0.91159 | 6.377 | 1.6021
| Rel.Std.Dev. |  | 1.286% | 1.95% | 1.694% | 0.3262% | 1.211% | 0.3342% | 1.278% | 0.84% | 1.367% | 2.227% | 0.6858% | 1.537% | 1.705% | 1.118% | 1.087% | 0.6014% | 0.4698% | 1.642% | 1.391% | 0.7993% | 5.213% | 1.28%
| PDM3 |  | 0.3293% | -9.152% | 1.374% | 0.5704% | -3.132% | -1.84% | 0.9319% | -2.026% | -2.711% | 0.3695% | 2.579% | 5.391% | 0.7043% | -4.27% | 0.9721% | 0.5704% | 4.735% | 0.9855% | -2.362% | -8.362% | -1.706% | 0.5704%
| Z-Score (Absolute) |  | 0.04 | 3.03 | 0.28 | 0.03 | 1.13 | 0.73 | 0.15 | 0.79 | 1 | 0.03 | 0.66 | 1.55 | 0.07 | 1.49 | 0.16 | 0.03 | 1.34 | 0.16 | 0.89 | 2.78 | 0.68 | 0.03
| Analytical results for P in OREAS 553b (Certified Value 0.035 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0349 | 0.036 | 0.047 | 0.037 | 0.031 | 0.0341 | 0.0327 | 0.035621 | 0.035 | 0.035 | 0.032 | 0.037 | 0.036 | 0.035717 | 0.0356 | 0.035 | 0.0358 | 0.033 | 0.031 | 0.036 | 0.0382 | 0.038 |  | 1
| 1 | 2 | 0.035 | 0.034 | 0.046 | 0.037 | 0.034 | 0.033 | 0.0344 | 0.03555 | 0.034 | 0.035 | 0.032 | 0.036 | 0.036 | 0.036572 | 0.0351 | 0.035 | 0.0377 | 0.0322 | 0.032 | 0.035 | 0.0405 | 0.037 |  | #N/A
| 1 | 3 | 0.035 | 0.035 | 0.045 | 0.036 | 0.033 | 0.0351 | 0.033 | 0.035161 | 0.035 | 0.034 | 0.033 | 0.035 | 0.036 | 0.037501 | 0.0352 | 0.036 | 0.0372 | 0.0323 | 0.032 | 0.035 | 0.0363 | 0.039 |  | 16
| 1 | 4 | 0.0353 | 0.035 | 0.046 | 0.037 | 0.032 | 0.0338 | 0.0344 | 0.035341 | 0.034 | 0.034 | 0.033 | 0.037 | 0.036 | 0.036458 | 0.0355 | 0.036 | 0.0374 | 0.032 | 0.032 | 0.035 | 0.0413 | 0.039 |  | 0.03505
| 1 | 5 | 0.0364 | 0.036 | 0.047 | 0.036 | 0.031 | 0.0323 | 0.0328 | 0.035244 | 0.035 | 0.034 | 0.034 | 0.037 | 0.036 | 0.036793 | 0.0356 | 0.036 | 0.0372 | 0.0318 | 0.032 | 0.035 | 0.0381 | 0.039 |  | 40
| 1 | 6 | 0.035 | 0.034 | 0.046 | 0.036 | 0.031 | 0.0322 | 0.0332 | 0.035541 | 0.035 | 0.035 | 0.033 | 0.037 | 0.036 | 0.036145 | 0.0357 | 0.035 | 0.041 | 0.0341 | 0.032 | 0.035 | 0.0382 | 0.037
| Mean |  | 0.035267 | 0.035 | 0.046167 | 0.0365 | 0.032 | 0.033417 | 0.033417 | 0.03541 | 0.034667 | 0.0345 | 0.032833 | 0.0365 | 0.036 | 0.036531 | 0.03545 | 0.0355 | 0.037717 | 0.032567 | 0.031833 | 0.035167 | 0.038767 | 0.038167
| Median |  | 0.035 | 0.035 | 0.046 | 0.0365 | 0.0315 | 0.0334 | 0.0331 | 0.035441 | 0.035 | 0.0345 | 0.033 | 0.037 | 0.036 | 0.036515 | 0.03555 | 0.0355 | 0.0373 | 0.03225 | 0.032 | 0.035 | 0.0382 | 0.0385
| Std Dev. |  | 0.00057155 | 0.00089443 | 0.00075277 | 0.00054772 | 0.0012649 | 0.0011268 | 0.00078081 | 0.00018761 | 0.0005164 | 0.00054772 | 0.00075277 | 0.00083666 | 0 | 0.0006049 | 0.0002429 | 0.00054772 | 0.0017371 | 0.00085479 | 0.00040825 | 0.00040825 | 0.0018217 | 0.00098319
| Rel.Std.Dev. |  | 1.621% | 2.556% | 1.631% | 1.501% | 3.953% | 3.372% | 2.337% | 0.5298% | 1.49% | 1.588% | 2.293% | 2.292% | 0% | 1.656% | 0.6852% | 1.543% | 4.606% | 2.625% | 1.282% | 1.161% | 4.699% | 2.576%
| PDM3 |  | 0.619% | -0.1418% | 31.72% | 4.138% | -8.701% | -4.659% | -4.659% | 1.027% | -1.093% | -1.568% | -6.324% | 4.138% | 2.711% | 4.226% | 1.142% | 1.285% | 7.609% | -7.084% | -9.177% | 0.3337% | 10.6% | 8.893%
| Z-Score (Absolute) |  | 0.04 | 0.19 | 6.2 | 0.67 | 1.91 | 1.1 | 1.1 | 0.04 | 0.38 | 0.48 | 1.43 | 0.67 | 0.38 | 0.68 | 0.06 | 0.09 | 1.36 | 1.59 | 2.01 | 0.1 | 1.96 | 1.62
| Analytical results for Pb in OREAS 553b (Certified Value 11.2 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 12.1 | 11.43 | 5 | 12 | 13.1 | 10.7 | 11.4114 | 10.7 | 11.5 | 10.9 | 12.1 | 10.5 | 12.8645 | 10 | 11.7 | 12.7 | 10 | 10.5 | 7 | 11 | 9.3 |  |  | 1
| 1 | 2 | 12.1 | 9.67 | 6 | 11 | 13.1 | 10.3 | 11.7508 | 10.7 | 12 | 10.8 | 11.7 | 10 | 12.298 | 10.3 | 11.9 | 12.6 | 10 | 10.7 | 6.5 | 11.1 | 9.1 |  |  | 13
| 1 | 3 | 11.7 | 10.62 | 6 | 11 | 13.8 | 10.9 | 11.985 | 11 | 11.1 | 11 | 11.5 | 10 | 12.1823 | 10 | 12.2 | 12.6 | 10 | 11 | 6.7 | 10.9 | 9 |  |  | 16
| 1 | 4 | 11.6 | 10.19 | 6 | 12 | 14 | 10.4 | 10.9117 | 10.7 | 11.2 | 11 | 11.6 | 9.9 | 12.5258 | 10.2 | 11.9 | 12.7 | 9 | 11 | 6.3 | 10.9 | 9.2 |  |  | 11.234
| 1 | 5 | 12.2 | 10.62 | 6 | 11 | 13.1 | 10.6 | 10.8352 | 10.4 | 11.5 | 11.2 | 11.8 | 10.2 | 12.4879 | 10.6 | 12.3 | 13.3 | 10 | 10.8 | 6.6 | 10.7 | 9.1 |  |  | 41
| 1 | 6 | 12 | 10.01 | 5 | 11 | 13.1 | 10.8 | 11.3788 | 11.2 | 11.6 | 11.1 | 12 | 9.7 | 11.8554 | 10.2 | 11.9 | 13.4 | 10 | 11.1 | 7.2 | 10.9 | 8.6
| Mean |  | 11.95 | 10.4233 | 5.6667 | 11.3333 | 13.3667 | 10.6167 | 11.3788 | 10.7833 | 11.4833 | 11 | 11.7833 | 10.05 | 12.369 | 10.2167 | 11.9833 | 12.8833 | 9.8333 | 10.85 | 6.7167 | 10.9167 | 9.05
| Median |  | 12.05 | 10.405 | 6 | 11 | 13.1 | 10.65 | 11.3951 | 10.7 | 11.5 | 11 | 11.75 | 10 | 12.3929 | 10.2 | 11.9 | 12.7 | 10 | 10.9 | 6.65 | 10.9 | 9.1
| Std Dev. |  | 0.2429 | 0.61376 | 0.5164 | 0.5164 | 0.41793 | 0.23166 | 0.45179 | 0.27869 | 0.31885 | 0.14142 | 0.23166 | 0.27386 | 0.34292 | 0.22286 | 0.22286 | 0.3656 | 0.40825 | 0.22583 | 0.33116 | 0.13292 | 0.2429
| Rel.Std.Dev. |  | 2.033% | 5.888% | 9.113% | 4.556% | 3.127% | 2.182% | 3.97% | 2.584% | 2.777% | 1.286% | 1.966% | 2.725% | 2.772% | 2.181% | 1.86% | 2.838% | 4.152% | 2.081% | 4.93% | 1.218% | 2.684%
| PDM3 |  | 6.374% | -7.216% | -49.56% | 0.8843% | 18.98% | -5.495% | 1.289% | -4.012% | 2.22% | -2.083% | 4.89% | -10.54% | 10.1% | -9.056% | 6.67% | 14.68% | -12.47% | -3.418% | -40.21% | -2.825% | -19.44%
| Z-Score (Absolute) |  | 0.85 | 0.46 | NA | NA | 2.07 | 0.29 | 0.36 | 0.15 | 0.45 | 0.04 | 0.71 | 0.78 | 1.21 | 0.64 | 0.88 | 1.66 | NA | 0.09 | 3.65 | 0.04 | 1.64
Results from laboratories 3, 4 and 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 553b (Certified Value 16.9 ppm) | [24 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 1
2
| 1 | 1 | 16.34 | 15 | 18.2 | 22.6968 | 16.87 | 17.4527 | 17.6 | 11.13 | [15 empty] | 1
| 1 | 2 | 16.66 | 14.3 | 18.4 | 22.5576 | 16.52 | 17.3794 | 17.2 | 11.1 | [15 empty] | 14
| 1 | 3 | 16.59 | 14.5 | 17.9 | 22.6232 | 16.33 | 17.9788 | 17.2 | 10.73 | [15 empty] | 16
| 1 | 4 | 17.02 | 14.5 | 18.3 | 22.5536 | 16.03 | 17.6832 | 19 | 11.19 | [15 empty] | 16.8542
| 1 | 5 | 17.13 | 15.5 | 18.3 | 22.644 | 16.58 | 17.3959 | 17 | 10.24 | [15 empty] | 42
| 1 | 6 | 17.08 | 14.7 | 17.9 | 22.5616 | 15.45 | 17.5601 | 17.2 | 8.85
| Mean |  | 16.8033 | 14.75 | 18.1667 | 22.6061 | 16.2967 | 17.575 | 17.5333 | 10.54
| Median |  | 16.84 | 14.6 | 18.25 | 22.5924 | 16.425 | 17.5064 | 17.2 | 10.915
| Std Dev. |  | 0.31967 | 0.43704 | 0.21602 | 0.058385 | 0.49959 | 0.22815 | 0.74476 | 0.90182
| Rel.Std.Dev. |  | 1.902% | 2.963% | 1.189% | 0.2583% | 3.066% | 1.298% | 4.248% | 8.556%
| PDM3 |  | -0.3016% | -12.48% | 7.787% | 34.13% | -3.308% | 4.277% | 4.03% | -37.46%
| Z-Score (Absolute) |  | 0.26 | 1.74 | 0.72 | 3.92 | 0.63 | 0.29 | 0.26 | 4.78
| Analytical results for Pt in OREAS 553b (Indicative Value 7.33 ppb) | [24 empty] | Indicative
Pt | Round | Replicate | Lab | [22 empty] | 1
| No | No | 25 | [22 empty] | ppb
|  |  | 4A*OES/MS | [22 empty] | 2
2
| 1 | 1 | 9 | [22 empty] | 1
| 1 | 2 | 8 | [22 empty] | 2
| 1 | 3 | 6 | [22 empty] | 16
| 1 | 4 | 8 | [22 empty] | 7.3333
| 1 | 5 | 6 | [22 empty] | 8
| 1 | 6 | 7
| Mean |  | 7.3333
| Median |  | 7.5
| Std Dev. |  | 1.2111
| Rel.Std.Dev. |  | 16.51%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 553b (Certified Value 96 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 94.93 | 61.1 | 93 | 90.8 | 133.093 | 93.3 | 103 | 93.1 | 93.2895 | 97 | 89.2 | 99.5 | 98.3 | 103 | 109.4825 | 91.7 | 91.4 | 99.05 | 75.4 | 67.6 | 90 | 95.5 | 123 | 1
| 1 | 2 | 95.3 | 59.8 | 96.6 | 91.4 | 134.833 | 89.8 | 101 | 95.2 | 97.5363 | 97.2 | 93 | 97.2 | 98.2 | 99.5 | 108.4654 | 94.1 | 95.6 | 100.08 | 74.7 | 69.8 | 96.9 | 93.3 | 132 | 15
| 1 | 3 | 95.59 | 63.6 | 81.9 | 91.4 | 133.2725 | 90.6 | 106 | 92.5 | 92.1861 | 98.3 | 84.4 | 98 | 95.1 | 101 | 106.2736 | 97.2 | 93.1 | 99.61 | 71.2 | 69.3 | 97.7 | 93.3 | 117 | 16
| 1 | 4 | 96.39 | 84.5 | 86.4 | 92.9 | 134.5245 | 90.4 | 104 | 93.9 | 91.1502 | 95.8 | 70.7 | 98.2 | 90.7 | 99.6 | 105.175 | 95 | 96.9 | 99.82 | 74.7 | 76 | 98.9 | 103 | 129 | 95.8242
| 1 | 5 | 97.2 | 53.8 | 86.6 | 92.3 | 133.6175 | 90.8 | 101 | 95.1 | 97.8886 | 93.7 | 92.3 | 98.8 | 90.9 | 102.5 | 100.2926 | 93.2 | 98.1 | 101.29 | 71.9 | 71.5 | 97 | 92.2 | 110 | 43
| 1 | 6 | 97.02 | 57 | 86.1 | 91 | 134.2885 | 89 | 102 | 92.4 | 91.8403 | 100.5 | 89.8 | 98 | 100.5 | 99 | 102.0484 | 94.3 | 96.1 | 102.06 | 73.9 | 81 | 97.9 | 90.2 | 92
| Mean |  | 96.0717 | 63.3 | 88.4333 | 91.6333 | 133.9382 | 90.65 | 102.8333 | 93.7 | 93.9818 | 97.0833 | 86.5667 | 98.2833 | 95.6167 | 100.7667 | 105.2896 | 94.25 | 95.2 | 100.3183 | 73.6333 | 72.5333 | 96.4 | 94.5833 | 117.1667
| Median |  | 95.99 | 60.45 | 86.5 | 91.4 | 133.953 | 90.5 | 102.5 | 93.5 | 92.7378 | 97.1 | 89.5 | 98.1 | 96.65 | 100.3 | 105.7243 | 94.2 | 95.85 | 99.95 | 74.3 | 70.65 | 97.35 | 93.3 | 120
| Std Dev. |  | 0.93858 | 10.923 | 5.3511 | 0.80664 | 0.71098 | 1.4529 | 1.9408 | 1.2442 | 2.9734 | 2.2921 | 8.3438 | 0.78592 | 4.1087 | 1.6813 | 3.5811 | 1.8382 | 2.4948 | 1.1308 | 1.6967 | 5.0405 | 3.2175 | 4.4692 | 14.689
| Rel.Std.Dev. |  | 0.977% | 17.26% | 6.051% | 0.8803% | 0.5308% | 1.603% | 1.887% | 1.328% | 3.164% | 2.361% | 9.639% | 0.7996% | 4.297% | 1.668% | 3.401% | 1.95% | 2.621% | 1.127% | 2.304% | 6.949% | 3.338% | 4.725% | 12.54%
| PDM3 |  | 0.2583% | -33.94% | -7.713% | -4.373% | 39.77% | -5.4% | 7.315% | -2.217% | -1.923% | 1.314% | -9.661% | 2.566% | -0.2165% | 5.158% | 9.878% | -1.643% | -0.6514% | 4.69% | -23.16% | -24.31% | 0.6009% | -1.295% | 22.27%
| Z-Score (Absolute) |  | 0.13 | 4.73 | 1 | 0.53 | 5.74 | 0.67 | 1.13 | 0.22 | 0.18 | 0.28 | 1.28 | 0.46 | 0.06 | 0.82 | 1.5 | 0.14 | 0 | 0.76 | 3.2 | 3.36 | 0.18 | 0.09 | 3.26
| Analytical results for Re in OREAS 553b (Certified Value 0.027 ppm) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 25 | 26 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.026 | 0.03 | 0.029 | < 0.05 | 0.034 | 0.06 | < 0.05 | 0.025 | 0.028 | 0.026 | 0.027 | 0.027 | < 0.01 | 0.024 | 0.004 | 0.027 | 0.027 | 0.0259 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 0.028 | 0.03 | 0.03 | < 0.05 | 0.032 | 0.028 | < 0.05 | 0.026 | 0.031 | 0.028 | 0.024 | 0.025 | < 0.01 | 0.029 | 0.003 | 0.026 | 0.026 | 0.0299 | < 0.1 |  |  |  |  | 16
| 1 | 3 | 0.027 | 0.023 | 0.027 | < 0.05 | 0.034 | 0.017 | < 0.05 | 0.028 | 0.026 | 0.025 | 0.029 | 0.025 | < 0.01 | 0.03 | 0.005 | 0.026 | 0.021 | 0.0283 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 0.028 | 0.03 | 0.03 | < 0.05 | 0.035 | 0.043 | < 0.05 | 0.029 | 0.025 | 0.024 | 0.028 | 0.028 | < 0.01 | 0.029 | < 0.003 | 0.028 | 0.029 | 0.0281 | < 0.1 |  |  |  |  | 0.027479
| 1 | 5 | 0.026 | 0.028 | 0.028 | < 0.05 | 0.039 | 0.038 | < 0.05 | 0.028 | 0.027 | 0.027 | 0.025 | 0.024 | < 0.01 | 0.033 | 0.005 | 0.027 | 0.029 | 0.0281 | < 0.1 |  |  |  |  | 44
| 1 | 6 | 0.029 | 0.031 | 0.029 | < 0.05 | 0.034 | 0.044 | < 0.05 | 0.03 | 0.032 | 0.026 | 0.03 | 0.026 | < 0.01 | 0.025 | 0.005 | 0.026 | 0.029 | 0.0292 | < 0.1
| Mean |  | 0.027333 | 0.028667 | 0.028833 |  | 0.034667 | 0.038333 |  | 0.027667 | 0.028167 | 0.026 | 0.027167 | 0.025833 |  | 0.028333 | 0.0044 | 0.026667 | 0.026833 | 0.02825
| Median |  | 0.0275 | 0.03 | 0.029 |  | 0.034 | 0.0405 |  | 0.028 | 0.0275 | 0.026 | 0.0275 | 0.0255 |  | 0.029 | 0.005 | 0.0265 | 0.028 | 0.0282
| Std Dev. |  | 0.0012111 | 0.0029439 | 0.001169 |  | 0.0023381 | 0.014733 |  | 0.0018619 | 0.0027869 | 0.0014142 | 0.0023166 | 0.001472 |  | 0.0033267 | 0.00089443 | 0.0008165 | 0.0031252 | 0.0013561
| Rel.Std.Dev. |  | 4.431% | 10.27% | 4.054% |  | 6.744% | 38.43% |  | 6.73% | 9.894% | 5.439% | 8.527% | 5.698% |  | 11.74% | 20.33% | 3.062% | 11.65% | 4.8%
| PDM3 |  | -0.5307% | 4.321% | 4.928% |  | 26.16% | 39.5% |  | 0.6823% | 2.502% | -5.383% | -1.137% | -5.989% |  | 3.108% | -83.99% | -2.957% | -2.35% | 2.805%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.76 | 1.18 | 3.71 | 5.56 | 1.18 | 0.17 | 0.42 | 0.67 | 0.08 | 0.76 | 11.29 | 0.51 | 11.84 | 0.34 | 0.25 | 0.46 | 11.46
| Analytical results for S in OREAS 553b (Certified Value 3.34 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 3.2864 | 3.3 | 3.38 | 3.45 | 3.3181 | 3.279 | 3.37 | 3.34 | 3.3324 | 3.27 | 3.38 | 3.29 | 3.44 | 3.34 | 3.3545 | 3.18 | 3.27 | 3.25 | 3.1385 | 3 | 3.35 | 3.45 | 3.632 | 1
| 1 | 2 | 3.2883 | 3.23 | 3.28 | 3.36 | 3.316 | 3.109 | 3.38 | 3.37 | 3.3574 | 3.25 | 3.42 | 3.2 | 3.4 | 3.32 | 3.4014 | 3.18 | 3.3 | 3.29 | 3.0818 | 3 | 3.36 | 3.49 | 3.597 | 31
| 1 | 3 | 3.3147 | 3.26 | 3.32 | 3.31 | 3.3165 | 2.868 | 3.39 | 3.46 | 3.3799 | 3.3 | 3.29 | 3.26 | 3.31 | 3.33 | 3.4184 | 3.25 | 3.2 | 3.29 | 3.1207 | 3 | 3.31 | 3.55 | 3.613 | 16
| 1 | 4 | 3.3522 | 3.26 | 3.33 | 3.4 | 3.3147 | 2.903 | 3.37 | 3.35 | 3.3637 | 3.28 | 3.35 | 3.28 | 3.43 | 3.32 | 3.453 | 3.23 | 3.31 | 3.29 | 2.9989 | 3 | 3.31 | 3.42 | 3.62 | 3.3376
| 1 | 5 | 3.3877 | 3.39 | 3.33 | 3.46 | 3.3182 | 3.023 | 3.37 | 3.48 | 3.3344 | 3.27 | 3.29 | 3.26 | 3.54 | 3.35 | 3.418 | 3.2 | 3.32 | 3.32 | 3.0371 | 3 | 3.32 | 3.46 | 3.567 | 45
| 1 | 6 | 3.3104 | 3.2 | 3.35 | 3.45 | 3.3133 | 2.906 | 3.38 | 3.31 | 3.3653 | 3.34 | 3.4 | 3.27 | 3.53 | 3.31 | 3.3298 | 3.19 | 3.31 | 3.34 | 3.1366 | 3 | 3.33 | 3.42 | 3.587
| Mean |  | 3.3233 | 3.2733 | 3.3317 | 3.405 | 3.3161 | 3.0147 | 3.3767 | 3.385 | 3.3555 | 3.285 | 3.355 | 3.26 | 3.4417 | 3.3283 | 3.3958 | 3.205 | 3.285 | 3.2967 | 3.0856 | 3 | 3.33 | 3.465 | 3.6027
| Median |  | 3.3125 | 3.26 | 3.33 | 3.425 | 3.3163 | 2.9645 | 3.375 | 3.36 | 3.3605 | 3.275 | 3.365 | 3.265 | 3.435 | 3.325 | 3.4097 | 3.195 | 3.305 | 3.29 | 3.1013 | 3 | 3.325 | 3.455 | 3.605
| Std Dev. |  | 0.039521 | 0.066232 | 0.033116 | 0.060249 | 0.0019169 | 0.15788 | 0.008165 | 0.06892 | 0.018654 | 0.031464 | 0.055408 | 0.031623 | 0.085654 | 0.01472 | 0.045543 | 0.02881 | 0.045056 | 0.030768 | 0.057472 | 0 | 0.020976 | 0.049295 | 0.023737
| Rel.Std.Dev. |  | 1.189% | 2.023% | 0.994% | 1.769% | 0.05781% | 5.237% | 0.2418% | 2.036% | 0.5559% | 0.9578% | 1.651% | 0.97% | 2.489% | 0.4423% | 1.341% | 0.8989% | 1.372% | 0.9333% | 1.863% | 0% | 0.6299% | 1.423% | 0.6589%
| PDM3 |  | -0.4285% | -1.925% | -0.1773% | 2.02% | -0.6427% | -9.675% | 1.171% | 1.421% | 0.537% | -1.576% | 0.5218% | -2.325% | 3.118% | -0.2772% | 1.746% | -3.972% | -1.576% | -1.226% | -7.55% | -10.11% | -0.2272% | 3.818% | 7.942%
| Z-Score (Absolute) |  | 0.06 | 0.67 | 0.04 | 0.94 | 0.15 | 3.85 | 0.59 | 0.69 | 0.33 | 0.53 | 0.33 | 0.84 | 1.39 | 0 | 0.83 | 1.51 | 0.53 | 0.39 | 2.98 | 4.03 | 0.02 | 1.68 | 3.36
| Analytical results for Sb in OREAS 553b (Certified Value 3.01 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 3.52 | 3.18 | 2.9 | 2.3 | < 5 | 3.23 | 1.234 | 3.2 | 3.14 | 3.21 | 3.15 | 3.03 | 3.3231 | 2.69 | 2.92 | 2.95 | 2.75 | 1.9 | 2.9 | 3.3 | 2.42 |  |  | 1
| 1 | 2 | 3.29 | 3.07 | 3.1 | 2.3 | < 5 | 3.31 | 1.6987 | 3.21 | 3.71 | 3.16 | 2.84 | 3.01 | 3.3625 | 2.59 | 2.71 | 3.03 | 2.48 | 2 | 3.2 | 3.2 | 2.36 |  |  | 17
| 1 | 3 | 3.44 | 3.11 | 2.2 | 2.4 | < 5 | 3.38 | 1.2318 | 3.28 | 3.04 | 3.12 | 2.63 | 2.97 | 3.3382 | 2.6 | 2.88 | 2.87 | 3.27 | 2.1 | 3.3 | 3.2 | 2.43 |  |  | 16
| 1 | 4 | 3.42 | 3.18 | 2.7 | 2.5 | < 5 | 3.3 | 1.1489 | 3.25 | 3.17 | 3.2 | 2.91 | 3.07 | 3.3903 | 2.52 | 3.24 | 2.53 | 1.46 | 2 | 3.1 | 3.2 | 2.62 |  |  | 3.0122
| 1 | 5 | 3.38 | 3.38 | 3.3 | 2.5 | < 5 | 3.23 | 1.1937 | 3.2 | 3.24 | 3.19 | 2.98 | 2.95 | 3.3053 | 2.58 | 3.05 | 2.13 | 3.15 | 1.9 | 3.2 | 3.1 | 2.52 |  |  | 46
| 1 | 6 | 3.43 | 3.08 | 3.6 | 2.3 | < 5 | 3.4 | 1.6332 | 3.46 | 3.33 | 3.36 | 2.94 | 3.05 | 3.3774 | 2.53 | 3.21 | 2.54 | 3.18 | 1.9 | 2.9 | 3.3 | 2.7
| Mean |  | 3.4133 | 3.1667 | 2.9667 | 2.3833 |  | 3.3083 | 1.3567 | 3.2667 | 3.2717 | 3.2067 | 2.9083 | 3.0133 | 3.3495 | 2.585 | 3.0017 | 2.675 | 2.715 | 1.9667 | 3.1 | 3.2167 | 2.5083
| Median |  | 3.425 | 3.145 | 3 | 2.35 |  | 3.305 | 1.2329 | 3.23 | 3.205 | 3.195 | 2.925 | 3.02 | 3.3504 | 2.585 | 2.985 | 2.705 | 2.95 | 1.95 | 3.15 | 3.2 | 2.475
| Std Dev. |  | 0.075807 | 0.11483 | 0.48854 | 0.098319 |  | 0.071949 | 0.24242 | 0.099933 | 0.23575 | 0.081894 | 0.17128 | 0.046332 | 0.032854 | 0.06091 | 0.20449 | 0.3394 | 0.68483 | 0.08165 | 0.16733 | 0.075277 | 0.1306
| Rel.Std.Dev. |  | 2.221% | 3.626% | 16.47% | 4.125% |  | 2.175% | 17.87% | 3.059% | 7.206% | 2.554% | 5.889% | 1.538% | 0.9809% | 2.356% | 6.813% | 12.69% | 25.22% | 4.152% | 5.398% | 2.34% | 5.207%
| PDM3 |  | 13.32% | 5.128% | -1.511% | -20.88% |  | 9.831% | -54.96% | 8.448% | 8.614% | 6.456% | -3.448% | 0.03785% | 11.2% | -14.18% | -0.3495% | -11.19% | -9.866% | -34.71% | 2.915% | 6.788% | -16.73%
| Z-Score (Absolute) |  | 0.97 | 0.39 | 0.08 | 1.45 | 1.18 | 0.72 | 3.87 | 0.62 | 0.64 | 0.48 | 0.22 | 0.03 | 0.82 | 0.98 | 0 | 0.77 | 0.67 | 2.43 | 0.23 | 0.51 | 1.16
| Analytical results for Sc in OREAS 553b (Certified Value 12.5 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 12.7 | 12 | 13 | 14 | 21.582 | 11.3 | 15.3 | 12.2 | 12.4296 | 11.8 | 12.2 | 12 | 13.3 | 13.4 | 13.2549 | 13.3 | 12.6 | 13.1 | 11.99 | 11 | 13.1 | 12 | 13.5 | 1
| 1 | 2 | 12.6 | 12 | 14 | 14 | 21.5239 | 11.2 | 14.9 | 11.8 | 12.8749 | 11.4 | 13.2 | 12 | 13 | 13.2 | 13.2638 | 13.1 | 12.5 | 13.1 | 11.75 | 11 | 13.1 | 12 | 13.2 | 18
| 1 | 3 | 12.5 | 12.1 | 13 | 14 | 21.4065 | 11.4 | 14.6 | 12.6 | 12.8628 | 11.3 | 12.2 | 11.9 | 13.1 | 13.4 | 13.4236 | 13.2 | 12.5 | 13 | 11.9 | 11 | 13.3 | 12 | 13.4 | 16
| 1 | 4 | 12.7 | 11.9 | 13 | 14 | 21.5024 | 11.2 | 15 | 12.3 | 11.9022 | 11.6 | 12.2 | 11.9 | 13.3 | 13.4 | 13.1165 | 13 | 12.9 | 13.2 | 11.56 | 11 | 13.1 | 13 | 13.6 | 12.5054
| 1 | 5 | 13 | 13.1 | 14 | 14 | 21.4725 | 11.7 | 13.9 | 11.8 | 11.7093 | 11.5 | 12.6 | 11.9 | 13.4 | 13.8 | 13.5217 | 13.3 | 12.3 | 13.6 | 11.48 | 11 | 13.1 | 11 | 12.8 | 47
| 1 | 6 | 12.7 | 12.4 | 13 | 13 | 21.6173 | 11.9 | 15.1 | 12.5 | 12.0774 | 12.4 | 12.8 | 11.8 | 13.5 | 12.7 | 13.2254 | 13.2 | 13.1 | 13.5 | 11.94 | 12 | 12.8 | 12 | 12
| Mean |  | 12.7 | 12.25 | 13.3333 | 13.8333 | 21.5175 | 11.45 | 14.8 | 12.2 | 12.3094 | 11.6667 | 12.5333 | 11.9167 | 13.2667 | 13.3167 | 13.301 | 13.1833 | 12.65 | 13.25 | 11.77 | 11.1667 | 13.0833 | 12 | 13.0833
| Median |  | 12.7 | 12.05 | 13 | 14 | 21.5132 | 11.35 | 14.95 | 12.25 | 12.2535 | 11.55 | 12.4 | 11.9 | 13.3 | 13.4 | 13.2594 | 13.2 | 12.55 | 13.15 | 11.825 | 11 | 13.1 | 12 | 13.3
| Std Dev. |  | 0.16733 | 0.45056 | 0.5164 | 0.40825 | 0.075817 | 0.2881 | 0.498 | 0.34059 | 0.49403 | 0.39833 | 0.41312 | 0.075277 | 0.18619 | 0.36009 | 0.14628 | 0.1169 | 0.29496 | 0.2429 | 0.21109 | 0.40825 | 0.16021 | 0.63246 | 0.60139
| Rel.Std.Dev. |  | 1.318% | 3.678% | 3.873% | 2.951% | 0.3524% | 2.516% | 3.365% | 2.792% | 4.013% | 3.414% | 3.296% | 0.6317% | 1.403% | 2.704% | 1.1% | 0.8868% | 2.332% | 1.833% | 1.793% | 3.656% | 1.225% | 5.27% | 4.597%
| PDM3 |  | 1.556% | -2.042% | 6.621% | 10.62% | 72.07% | -8.439% | 18.35% | -2.442% | -1.568% | -6.707% | 0.2234% | -4.708% | 6.088% | 6.487% | 6.362% | 5.421% | 1.156% | 5.954% | -5.881% | -10.71% | 4.622% | -4.041% | 4.622%
| Z-Score (Absolute) |  | 0.03 | 0.49 | NA | NA | 10.22 | 1.42 | 2.46 | 0.55 | 0.42 | 1.17 | 0.16 | 0.88 | 0.68 | 0.74 | 0.72 | 0.59 | 0.03 | 0.66 | 1.05 | 1.74 | 0.47 | NA | 0.47
Results from laboratories 3, 4 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 553b (Certified Value 5.73 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 5.2 | 5.9 | 5.6 | 6 | 7 | 7 | 5 | 5.1408 | 6 | 6 | 6 | 6 | 6 | 5.807 | 5.7 | 6 | 5.5 | 4.32 | 5 | 5.54 | < 10 |  |  | 1
| 1 | 2 | 5.4 | 5.4 | 5.7 | 6 | 7 | 7 | 6 | 5.8063 | 5 | 7 | 6 | 5 | 6 | 6.3862 | 5.47 | 6 | 6 | 4.09 | 5 | 4.9 | < 10 |  |  | 21
| 1 | 3 | 5.3 | 5.8 | 5.2 | 6 | 7 | 6 | 5 | 5.7555 | 5 | 5 | 6 | 6 | 6 | 6.1894 | 5.32 | 6 | 6.3 | 4.85 | 5 | 5.67 | < 10 |  |  | 16
| 1 | 4 | 5.3 | 6.6 | 5.7 | 6 | 7 | 6 | 7 | 5.2319 | 6 | 6 | 6 | 6 | 6 | 6.2167 | 5.78 | 6 | 5.7 | 1.31 | 5 | 5.16 | < 10 |  |  | 5.729
| 1 | 5 | 5.6 | 6.4 | 5.7 | 6 | 7 | 6 | 5 | 5.2807 | 5 | 7 | 6 | 6 | 6 | 5.8491 | 5.59 | 6 | 5.7 | 4.74 | 5 | 4.34 | < 10 |  |  | 48
| 1 | 6 | 5.6 | 4.8 | 5.7 | 6 | 7 | 6 | 6 | 5.246 | 6 | 6 | 6 | 6 | 5 | 6.4174 | 5.68 | 6 | 5.6 | 4.91 | 5 | 4.55 | < 10
| Mean |  | 5.4 | 5.8167 | 5.6 | 6 | 7 | 6.3333 | 5.6667 | 5.4102 | 5.5 | 6.1667 | 6 | 5.8333 | 5.8333 | 6.1443 | 5.59 | 6 | 5.8 | 4.0367 | 5 | 5.0267
| Median |  | 5.35 | 5.85 | 5.7 | 6 | 7 | 6 | 5.5 | 5.2633 | 5.5 | 6 | 6 | 6 | 6 | 6.203 | 5.635 | 6 | 5.7 | 4.53 | 5 | 5.03
| Std Dev. |  | 0.16733 | 0.65853 | 0.2 | 0 | 0 | 0.5164 | 0.8165 | 0.2913 | 0.54772 | 0.75277 | 0 | 0.40825 | 0.40825 | 0.26126 | 0.16947 | 0 | 0.29665 | 1.3738 | 0 | 0.53102
| Rel.Std.Dev. |  | 3.099% | 11.32% | 3.571% | 0% | 0% | 8.154% | 14.41% | 5.384% | 9.959% | 12.21% | 0% | 6.999% | 6.999% | 4.252% | 3.032% | 0% | 5.115% | 34.03% | 0% | 10.56%
| PDM3 |  | -5.742% | 1.531% | -2.251% | 4.731% | 22.19% | 10.55% | -1.087% | -5.564% | -3.996% | 7.64% | 4.731% | 1.822% | 1.822% | 7.25% | -2.425% | 4.731% | 1.24% | -29.54% | -12.72% | -12.26%
| Z-Score (Absolute) |  | 0.9 | 0.04 | 0.45 | 0.45 | 2.7 | 1.2 | 0.3 | 0.88 | 0.67 | 0.82 | 0.45 | 0.07 | 0.07 | 0.77 | 0.47 | 0.45 | 0 | 3.96 | 1.8 | 1.74 | 1.8
| Analytical results for Sm in OREAS 553b (Certified Value 8.69 ppm) | [24 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 8.41 | 7.3 | 9.9 | 10.284 | 7.8 | 8.9153 | 8.63 | 5.8 | [15 empty] | 1
| 1 | 2 | 8.55 | 7.4 | 9.7 | 10.2576 | 7.83 | 8.4729 | 8.55 | 6 | [15 empty] | 5
| 1 | 3 | 8.74 | 7.7 | 8.8 | 10.2016 | 7.92 | 9.083 | 8.67 | 5.7 | [15 empty] | 16
| 1 | 4 | 8.67 | 7.4 | 8.4 | 10.2264 | 7.78 | 8.7765 | 8.36 | 6 | [15 empty] | 8.6857
| 1 | 5 | 8.83 | 7.8 | 9.3 | 10.2712 | 7.9 | 8.672 | 8.18 | 5.6 | [15 empty] | 49
| 1 | 6 | 8.91 | 7.3 | 9.7 | 10.1952 | 7.7 | 8.8558 | 8.46 | 4.8
| Mean |  | 8.685 | 7.4833 | 9.3 | 10.2393 | 7.8217 | 8.7959 | 8.475 | 5.65
| Median |  | 8.705 | 7.4 | 9.5 | 10.242 | 7.815 | 8.8161 | 8.505 | 5.75
| Std Dev. |  | 0.18371 | 0.2137 | 0.58992 | 0.037103 | 0.081097 | 0.2099 | 0.18338 | 0.44609
| Rel.Std.Dev. |  | 2.115% | 2.856% | 6.343% | 0.3624% | 1.037% | 2.386% | 2.164% | 7.895%
| PDM3 |  | -0.008609% | -13.84% | 7.072% | 17.89% | -9.948% | 1.268% | -2.426% | -34.95%
| Z-Score (Absolute) |  | 0.1 | 1 | 0.66 | 1.51 | 0.69 | 0.2 | 0.1 | 2.67
| Analytical results for Sn in OREAS 553b (Certified Value 2.81 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.9 | 2.7 | 3 | 3.1 | 2.7 | 2.8 | 2.6371 | 2.9 | 2.7 | 2.8 | 2.8 | 2.7 | 2.8916 | 2.7 | 2.6 | 2.8 | 3.48 | 2.3 | 2.8 | 2.9 | 2.3 |  |  | 1
| 1 | 2 | 2.9 | 2.5 | 3 | 3.1 | 3.2 | 2.6 | 2.8312 | 2.9 | 2.8 | 2.7 | 2.7 | 2.6 | 2.9315 | 2.67 | 2.7 | 2.8 | 2.38 | 2.3 | 2.8 | 2.8 | 2.3 |  |  | 22
| 1 | 3 | 2.8 | 2.6 | 3 | 2.9 | 3 | 2.8 | 2.8139 | 2.9 | 2.6 | 2.9 | 2.7 | 2.6 | 2.9643 | 2.59 | 2.6 | 2.9 | 2.98 | 2.4 | 3 | 2.8 | 2.3 |  |  | 16
| 1 | 4 | 2.9 | 2.6 | 3 | 2.8 | 3 | 2.6 | 2.6951 | 2.8 | 2.8 | 2.8 | 2.7 | 2.6 | 2.986 | 2.64 | 2.8 | 2.8 | 2.99 | 2.4 | 3.1 | 3 | 2.4 |  |  | 2.8123
| 1 | 5 | 3 | 2.8 | 3 | 3 | 2.8 | 2.8 | 2.6063 | 2.9 | 2.8 | 2.8 | 2.8 | 2.5 | 2.863 | 2.77 | 2.8 | 2.8 | 2.78 | 2.3 | 2.9 | 2.7 | 2.3 |  |  | 50
| 1 | 6 | 2.9 | 2.6 | 3 | 2.9 | 3.3 | 2.7 | 2.6613 | 3.1 | 2.8 | 2.9 | 2.7 | 2.5 | 2.9976 | 2.86 | 2.7 | 2.9 | 2.78 | 2.4 | 2.9 | 2.8 | 2.3
| Mean |  | 2.9 | 2.6333 | 3 | 2.9667 | 3 | 2.7167 | 2.7075 | 2.9167 | 2.75 | 2.8167 | 2.7333 | 2.5833 | 2.939 | 2.705 | 2.7 | 2.8333 | 2.8983 | 2.35 | 2.9167 | 2.8333 | 2.3167
| Median |  | 2.9 | 2.6 | 3 | 2.95 | 3 | 2.75 | 2.6782 | 2.9 | 2.8 | 2.8 | 2.7 | 2.6 | 2.9479 | 2.685 | 2.7 | 2.8 | 2.88 | 2.35 | 2.9 | 2.8 | 2.3
| Std Dev. |  | 0.063246 | 0.10328 | 0 | 0.12111 | 0.22804 | 0.098319 | 0.093931 | 0.098319 | 0.083666 | 0.075277 | 0.05164 | 0.075277 | 0.053571 | 0.096902 | 0.089443 | 0.05164 | 0.36058 | 0.054772 | 0.1169 | 0.10328 | 0.040825
| Rel.Std.Dev. |  | 2.181% | 3.922% | 0% | 4.082% | 7.601% | 3.619% | 3.469% | 3.371% | 3.042% | 2.673% | 1.889% | 2.914% | 1.823% | 3.582% | 3.313% | 1.823% | 12.44% | 2.331% | 4.008% | 3.645% | 1.762%
| PDM3 |  | 3.119% | -6.363% | 6.675% | 5.489% | 6.675% | -3.4% | -3.726% | 3.711% | -2.215% | 0.1556% | -2.808% | -8.141% | 4.506% | -3.815% | -3.993% | 0.7483% | 3.06% | -16.44% | 3.711% | 0.7483% | -17.62%
| Z-Score (Absolute) |  | 0.51 | 1.13 | 1.13 | 0.93 | 1.13 | 0.62 | 0.67 | 0.62 | 0.41 | 0 | 0.51 | 1.44 | 0.76 | 0.69 | 0.72 | 0.1 | 0.5 | 2.88 | 0.62 | 0.1 | 3.09
| Analytical results for Sr in OREAS 553b (Certified Value 70 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 70.36 | 68 | 70.6 | 69.3 | 69.6375 | 83.4 | 68 | 63.4 | 70.6451 | 71.2 | 71.5 | 68 | 71.8 | 72.6 | 70.0253 | 67.2 | 67.2 | 72.46 | 76 | 58 | 68 | 60.5 | 62 | 1
| 1 | 2 | 69.62 | 69 | 71.2 | 68.6 | 70.5405 | 79 | 67 | 67.1 | 74.8085 | 70.6 | 75.3 | 67 | 71.3 | 68.8 | 70.9822 | 67.7 | 67.3 | 72.72 | 76 | 59 | 68 | 63.1 | 63 | #N/A
| 1 | 3 | 71.01 | 68 | 67.2 | 68.4 | 68.9445 | 82.5 | 71 | 65.9 | 75.7669 | 71.5 | 68 | 70.1 | 71.8 | 71.3 | 71.4211 | 66.1 | 67 | 71.67 | 76 | 58 | 68 | 63.1 | 61 | 16
| 1 | 4 | 71.37 | 67 | 70.9 | 68.6 | 68.4825 | 74.6 | 68 | 63.7 | 70.1249 | 70.1 | 69.5 | 67.5 | 72.9 | 70.8 | 71.5296 | 66.7 | 67.9 | 72.8 | 74 | 61 | 69 | 61.2 | 62 | 70.0375
| 1 | 5 | 71.59 | 72 | 71.5 | 68 | 68.22 | 66.8 | 66 | 67.5 | 70.4117 | 69.1 | 73.2 | 67.7 | 72.8 | 70.3 | 71.0236 | 67.4 | 67.8 | 72.79 | 75 | 58 | 68 | 58.6 | 58 | 51
| 1 | 6 | 70.78 | 68 | 71 | 68.1 | 67.779 | 68 | 68 | 65.6 | 72.3979 | 73.4 | 74 | 66.6 | 74.3 | 70.4 | 69.2334 | 68 | 67.3 | 72.16 | 77 | 61 | 68 | 60.8 | 52
| Mean |  | 70.7883 | 68.6667 | 70.4 | 68.5 | 68.934 | 75.7167 | 68 | 65.5333 | 72.3592 | 70.9833 | 71.9167 | 67.8167 | 72.4833 | 70.7 | 70.7025 | 67.1833 | 67.4167 | 72.4333 | 75.6667 | 59.1667 | 68.1667 | 61.2167 | 59.6667
| Median |  | 70.895 | 68 | 70.95 | 68.5 | 68.7135 | 76.8 | 68 | 65.75 | 71.5215 | 70.9 | 72.35 | 67.6 | 72.3 | 70.6 | 71.0029 | 67.3 | 67.3 | 72.59 | 76 | 58.5 | 68 | 61 | 61.5
| Std Dev. |  | 0.71815 | 1.7512 | 1.5962 | 0.46476 | 1.0123 | 7.1572 | 1.6733 | 1.6955 | 2.4222 | 1.458 | 2.7853 | 1.2254 | 1.087 | 1.2522 | 0.89429 | 0.69113 | 0.35449 | 0.44765 | 1.0328 | 1.472 | 0.40825 | 1.7105 | 4.1312
| Rel.Std.Dev. |  | 1.015% | 2.55% | 2.267% | 0.6785% | 1.468% | 9.453% | 2.461% | 2.587% | 3.348% | 2.054% | 3.873% | 1.807% | 1.5% | 1.771% | 1.265% | 1.029% | 0.5258% | 0.618% | 1.365% | 2.488% | 0.5989% | 2.794% | 6.924%
| PDM3 |  | 1.072% | -1.957% | 0.5175% | -2.195% | -1.576% | 8.109% | -2.909% | -6.431% | 3.315% | 1.35% | 2.683% | -3.171% | 3.492% | 0.9459% | 0.9495% | -4.075% | -3.742% | 3.421% | 8.037% | -15.52% | -2.671% | -12.59% | -14.81%
| Z-Score (Absolute) |  | 0.67 | 0.1 | 0.53 | 0.16 | 0 | 2.47 | 0.34 | 1.24 | 1.25 | 0.75 | 1.08 | 0.41 | 1.29 | 0.64 | 0.64 | 0.64 | 0.55 | 1.27 | 2.45 | 3.55 | 0.28 | 2.81 | 3.37
| Analytical results for Ta in OREAS 553b (Certified Value 1.26 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 1.37 | 1.5 | 1.9 | 0.64 | 1.44 | 0.51 | 1.69 | 0.93005 | 1.28 | 1.36 | 1.31 | 1.13 | 1.37 | 1.3329 | 0.85 | 1.18 | 0.23 | 0.7 | 1.2 | 1.26 | 1.2 |  |  | 1
| 1 | 2 | 1.45 | 1.3 | 1.9 | 0.68 | 1.58 | 0.42 | 1.51 | 0.87976 | 1.27 | 1.47 | 1.3 | 1.15 | 1.36 | 1.2478 | 0.85 | 1.16 | 0.21 | 0.7 | 1.1 | 1.47 | 1.15 |  |  | 7
| 1 | 3 | 1.44 | 1.3 | 1.8 | 0.61 | 1.47 | 0.51 | 1.64 | 0.90399 | 1.3 | 1.4 | 1.3 | 1.1 | 1.33 | 1.3407 | 0.88 | 1.13 | 0.3 | 0.7 | 1.1 | 1.36 | 1.29 |  |  | 16
| 1 | 4 | 1.45 | 1.4 | 1.9 | 0.56 | 1.48 | 0.46 | 1.58 | 0.89876 | 1.31 | 1.33 | 1.33 | 1.12 | 1.36 | 1.316 | 0.8 | 1.12 | 0.12 | 0.7 | 1.1 | 1.14 | 1.22 |  |  | 1.2644
| 1 | 5 | 1.26 | 1.2 | 1.9 | 0.68 | 1.62 | 0.43 | 1.5 | 0.91783 | 1.26 | 1.36 | 1.35 | 1.12 | 1.35 | 1.3024 | 0.92 | 1.16 | 0.31 | 0.7 | 1 | 1.4 | 1.32 |  |  | 52
| 1 | 6 | 1.6 | 1.3 | 1.9 | 0.64 | 1.53 | 0.47 | 1.65 | 0.88139 | 1.34 | 1.37 | 1.3 | 1.13 | 1.37 | 1.3097 | 0.8 | 1.04 | 0.31 | 0.7 | 1.1 | 1.35 | 1.38
| Mean |  | 1.4283 | 1.3333 | 1.8833 | 0.635 | 1.52 | 0.46667 | 1.595 | 0.90196 | 1.2933 | 1.3817 | 1.315 | 1.125 | 1.3567 | 1.3083 | 0.85 | 1.1317 | 0.24667 | 0.7 | 1.1 | 1.33 | 1.26
| Median |  | 1.445 | 1.3 | 1.9 | 0.64 | 1.505 | 0.465 | 1.61 | 0.90138 | 1.29 | 1.365 | 1.305 | 1.125 | 1.36 | 1.3129 | 0.85 | 1.145 | 0.265 | 0.7 | 1.1 | 1.355 | 1.255
| Std Dev. |  | 0.11161 | 0.10328 | 0.040825 | 0.045497 | 0.06957 | 0.038297 | 0.078166 | 0.019854 | 0.029439 | 0.048751 | 0.020736 | 0.016432 | 0.015055 | 0.03289 | 0.046476 | 0.049967 | 0.075542 | 1.2162e-16 | 0.063246 | 0.11559 | 0.085088
| Rel.Std.Dev. |  | 7.814% | 7.746% | 2.168% | 7.165% | 4.577% | 8.207% | 4.901% | 2.201% | 2.276% | 3.528% | 1.577% | 1.461% | 1.11% | 2.514% | 5.468% | 4.415% | 30.63% | 1.737e-14% | 5.75% | 8.691% | 6.753%
| PDM3 |  | 12.97% | 5.453% | 48.95% | -49.78% | 20.22% | -63.09% | 26.15% | -28.66% | 2.289% | 9.275% | 4.003% | -11.02% | 7.298% | 3.47% | -32.77% | -10.5% | -80.49% | -44.64% | -13% | 5.189% | -0.3472%
| Z-Score (Absolute) |  | 0.54 | 0.16 | 2.36 | 2.64 | 0.91 | 3.31 | 1.21 | 1.57 | 0 | 0.35 | 0.09 | 0.67 | 0.25 | 0.06 | 1.78 | 0.65 | 4.19 | 2.38 | 0.77 | 0.15 | 0.13
| Analytical results for Tb in OREAS 553b (Certified Value 0.74 ppm) | [24 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 15 | 16 | 17 | 25 | 26 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 0.69 | 0.7 | 0.9 | 1.1259 | 0.86 | 0.67 | 0.79092 | 0.81 | 0.69 | 0.76 | 0.57 | [12 empty] | 1
| 1 | 2 | 0.68 | 0.7 | 1 | 1.1264 | 0.88 | 0.71 | 0.82722 | 0.86 | 0.73 | 0.77 | 0.55 | [12 empty] | 8
| 1 | 3 | 0.69 | 0.7 | 1 | 1.1225 | 0.82 | 0.69 | 0.80051 | 0.84 | 0.72 | 0.76 | 0.55 | [12 empty] | 16
| 1 | 4 | 0.74 | 0.7 | 1 | 1.1717 | 0.81 | 0.73 | 0.78922 | 0.84 | 0.74 | 0.73 | 0.58 | [12 empty] | 0.74168
| 1 | 5 | 0.72 | 0.7 | 1 | 1.173 | 0.85 | 0.68 | 0.77073 | 0.85 | 0.74 | 0.74 | 0.53 | [12 empty] | 53
| 1 | 6 | 0.71 | 0.7 | 1 | 1.15 | 0.88 | 0.71 | 0.77592 | 0.86 | 0.76 | 0.79 | 0.47
| Mean |  | 0.705 | 0.7 | 0.98333 | 1.1449 | 0.85 | 0.69833 | 0.79242 | 0.84333 | 0.73 | 0.75833 | 0.54167
| Median |  | 0.7 | 0.7 | 1 | 1.1382 | 0.855 | 0.7 | 0.79007 | 0.845 | 0.735 | 0.76 | 0.55
| Std Dev. |  | 0.022583 | 1.2162e-16 | 0.040825 | 0.02341 | 0.029665 | 0.022286 | 0.02016 | 0.018619 | 0.023664 | 0.02137 | 0.0392
| Rel.Std.Dev. |  | 3.203% | 1.737e-14% | 4.152% | 2.045% | 3.49% | 3.191% | 2.544% | 2.208% | 3.242% | 2.818% | 7.237%
| PDM3 |  | -4.945% | -5.619% | 32.58% | 54.37% | 14.61% | -5.844% | 6.842% | 13.71% | -1.574% | 2.246% | -26.97%
| Z-Score (Absolute) |  | 0.6 | NA | NA | 4.34 | 1.03 | 0.67 | 0.38 | 0.96 | 0.32 | 0 | 2.44
Results from laboratories 2 and 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 553b (Certified Value 0.063 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 25 | 26 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | 0.1 | < 0.1 | < 0.1 | 0.49 | 0.06 | < 0.1 | 0.1 | 0.05 | 0.08 | 0.05 | 0.07 | < 0.2 | 0.06 | < 0.03 | < 0.5 | < 0.1 | 0.08 | < 0.05 | [5 empty] | 1
| 1 | 2 | 0.06 | < 0.1 | < 0.1 | 0.45 | 0.07 | < 0.1 | 0.07 | 0.05 | 0.06 | 0.05 | 0.07 | < 0.2 | 0.06 | < 0.03 | < 0.5 | 0.1 | 0.08 | 0.05 | [5 empty] | 23
| 1 | 3 | < 0.05 | < 0.1 | < 0.1 | 0.48 | 0.06 | < 0.1 | 0.07 | 0.05 | 0.07 | 0.06 | 0.06 | < 0.2 | 0.06 | < 0.03 | < 0.5 | < 0.1 | 0.09 | < 0.05 | [5 empty] | 16
| 1 | 4 | 0.06 | < 0.1 | < 0.1 | 0.41 | 0.08 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.06 | 0.07 | < 0.2 | < 0.05 | < 0.03 | < 0.5 | < 0.1 | 0.08 | < 0.05 | [5 empty] | 0.063233
| 1 | 5 | 0.06 | < 0.1 | < 0.1 | 0.48 | 0.05 | < 0.1 | 0.07 | 0.07 | 0.07 | < 0.05 | 0.06 | < 0.2 | 0.06 | < 0.03 | < 0.5 | < 0.1 | 0.08 | < 0.05 | [5 empty] | 54
| 1 | 6 | 0.06 | < 0.1 | < 0.1 | 0.47 | 0.08 | < 0.1 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | < 0.2 | 0.06 | < 0.03 | < 0.5 | < 0.1 | 0.09 | < 0.05
| Mean |  | 0.068 |  |  | 0.46333 | 0.066667 |  | 0.073333 | 0.058333 | 0.066667 | 0.056 | 0.063333 |  | 0.06 |  |  | 0.1 | 0.083333 | 0.05
| Median |  | 0.06 |  |  | 0.475 | 0.065 |  | 0.07 | 0.055 | 0.07 | 0.06 | 0.065 |  | 0.06 |  |  | 0.1 | 0.08 | 0.05
| Std Dev. |  | 0.017889 |  |  | 0.029439 | 0.012111 |  | 0.013663 | 0.0098319 | 0.010328 | 0.0054772 | 0.008165 |  | 0 |  |  |  | 0.005164
| Rel.Std.Dev. |  | 26.31% |  |  | 6.354% | 18.17% |  | 18.63% | 16.85% | 15.49% | 9.781% | 12.89% |  | 0% |  |  |  | 6.197%
| PDM3 |  | 7.538% |  |  | 632.7% | 5.43% |  | 15.97% | -7.749% | 5.43% | -11.44% | 0.1581% |  | -5.113% |  |  | 58.14% | 31.79% | -20.93%
| Z-Score (Absolute) |  | 0.09 | 0.67 | 0.67 | 28.41 | 0.5 | 0.67 | 0.97 | 0.09 | 0.5 | 0.62 | 0.26 | 2.84 | 0.38 | 3.14 | 13.4 | 0.09 | 1.67 | 2.14
| Analytical results for Th in OREAS 553b (Certified Value 15.5 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 15.19 | 14.3 | 15.6 | 14.6 | 15.18 | 15.71 | 16 | 15.5378 | 16.5 | 15.65 | 15.4 | 15.15 | 15.2 | 16.1224 | 15.85 | 15.08 | 15.74 | 15.12 | 12.2 | 14.2 | 16.1 | 15.19 |  | 1
| 1 | 2 | 14.94 | 14 | 16.7 | 15 | 15.32 | 15.12 | 14.9 | 14.8004 | 16.6 | 16.55 | 15.1 | 15.1 | 14.8 | 15.5865 | 16.14 | 14.6 | 15.6 | 14.12 | 12.8 | 15.1 | 16.3 | 15.48 |  | 10
| 1 | 3 | 15.41 | 14.8 | 15.6 | 14.4 | 15.55 | 15.86 | 15.9 | 14.9832 | 17 | 15.65 | 15.4 | 14.85 | 15.25 | 16.2841 | 16.04 | 15.78 | 15.26 | 17.17 | 12.9 | 15 | 16 | 15.69 |  | 16
| 1 | 4 | 15.21 | 14.3 | 15.8 | 14.9 | 15.27 | 15.65 | 15 | 15.006 | 15.95 | 14.3 | 15.4 | 15.05 | 14.65 | 16.2665 | 15.9 | 15.06 | 15.29 | 11.32 | 13.8 | 15.3 | 15.3 | 15.81 |  | 15.4502
| 1 | 5 | 15.47 | 15.8 | 16.1 | 14.5 | 15.89 | 14.79 | 15.8 | 14.934 | 15.5 | 15.85 | 15.55 | 14.5 | 15.1 | 15.9722 | 16.19 | 15.6 | 15.57 | 16.94 | 13.9 | 14.8 | 15.9 | 15.65 |  | 55
| 1 | 6 | 15.3 | 14.8 | 15.4 | 14.5 | 15.64 | 15.56 | 15.4 | 15.6427 | 17.7 | 15.9 | 15.45 | 15.45 | 14.65 | 15.9384 | 15.93 | 15.44 | 15.3 | 16.67 | 13.9 | 15 | 15.9 | 12.61
| Mean |  | 15.2533 | 14.6667 | 15.8667 | 14.65 | 15.475 | 15.4483 | 15.5 | 15.1507 | 16.5417 | 15.65 | 15.3833 | 15.0167 | 14.9417 | 16.0283 | 16.0083 | 15.26 | 15.46 | 15.2233 | 13.25 | 14.9 | 15.9167 | 15.0717
| Median |  | 15.255 | 14.55 | 15.7 | 14.55 | 15.435 | 15.605 | 15.6 | 14.9946 | 16.55 | 15.75 | 15.4 | 15.075 | 14.95 | 16.0473 | 15.985 | 15.26 | 15.435 | 15.895 | 13.35 | 15 | 15.95 | 15.565
| Std Dev. |  | 0.18854 | 0.6377 | 0.47188 | 0.2429 | 0.26764 | 0.40799 | 0.47329 | 0.34946 | 0.77357 | 0.73959 | 0.15055 | 0.31885 | 0.27462 | 0.2598 | 0.13732 | 0.43044 | 0.20229 | 2.2491 | 0.71764 | 0.37947 | 0.33714 | 1.2249
| Rel.Std.Dev. |  | 1.236% | 4.348% | 2.974% | 1.658% | 1.729% | 2.641% | 3.053% | 2.307% | 4.677% | 4.726% | 0.9787% | 2.123% | 1.838% | 1.621% | 0.8578% | 2.821% | 1.308% | 14.77% | 5.416% | 2.547% | 2.118% | 8.127%
| PDM3 |  | -1.274% | -5.071% | 2.696% | -5.179% | 0.1607% | -0.01192% | 0.3225% | -1.938% | 7.065% | 1.293% | -0.4326% | -2.806% | -3.291% | 3.742% | 3.613% | -1.231% | 0.06359% | -1.468% | -14.24% | -3.561% | 3.019% | -2.45%
| Z-Score (Absolute) |  | 0.15 | 1.39 | 1.16 | 1.43 | 0.33 | 0.27 | 0.38 | 0.36 | 2.6 | 0.7 | 0.13 | 0.65 | 0.81 | 1.5 | 1.46 | 0.13 | 0.29 | 0.21 | 4.41 | 0.9 | 1.27 | 0.53
| Analytical results for Ti in OREAS 553b (Certified Value 0.258 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.3007 | 0.354 | 0.484 | 0.146 | 0.2238 | 0.128 | 0.17 | 0.16 | 0.19231 | 0.308 | 0.341 | 0.316 | 0.275 | 0.344 | 0.30091 | 0.35 | 0.222 | 0.251 | 0.0993 | 0.18 | 0.256 | 0.269 | 0.302 | 1
| 1 | 2 | 0.3088 | 0.34 | 0.473 | 0.158 | 0.22344 | 0.129 | 0.16 | 0.15 | 0.19172 | 0.313 | 0.34 | 0.306 | 0.272 | 0.337 | 0.30434 | 0.35 | 0.221 | 0.2606 | 0.1024 | 0.186 | 0.243 | 0.314 | 0.287 | 26
| 1 | 3 | 0.3191 | 0.325 | 0.461 | 0.144 | 0.22347 | 0.133 | 0.18 | 0.14 | 0.18574 | 0.307 | 0.335 | 0.313 | 0.271 | 0.33 | 0.31057 | 0.35 | 0.233 | 0.2675 | 0.1035 | 0.187 | 0.251 | 0.306 | 0.311 | 16
| 1 | 4 | 0.3279 | 0.336 | 0.463 | 0.163 | 0.22376 | 0.13 | 0.16 | 0.15 | 0.18341 | 0.315 | 0.318 | 0.316 | 0.275 | 0.344 | 0.31369 | 0.34 | 0.228 | 0.2521 | 0.1037 | 0.198 | 0.253 | 0.257 | 0.302 | 0.25811
| 1 | 5 | 0.298 | 0.325 | 0.469 | 0.162 | 0.22376 | 0.125 | 0.15 | 0.16 | 0.19761 | 0.314 | 0.324 | 0.315 | 0.273 | 0.335 | 0.30991 | 0.38 | 0.226 | 0.2626 | 0.1027 | 0.185 | 0.248 | 0.29 | 0.318 | 56
| 1 | 6 | 0.3517 | 0.33 | 0.476 | 0.175 | 0.2235 | 0.13 | 0.17 | 0.14 | 0.19309 | 0.319 | 0.322 | 0.307 | 0.271 | 0.343 | 0.30136 | 0.37 | 0.225 | 0.2715 | 0.1036 | 0.19 | 0.24 | 0.287 | 0.347
| Mean |  | 0.3177 | 0.335 | 0.471 | 0.158 | 0.22362 | 0.12917 | 0.165 | 0.15 | 0.19065 | 0.31267 | 0.33 | 0.31217 | 0.27283 | 0.33883 | 0.3068 | 0.35667 | 0.22583 | 0.26088 | 0.10253 | 0.18767 | 0.2485 | 0.28717 | 0.31117
| Median |  | 0.31395 | 0.333 | 0.471 | 0.16 | 0.22363 | 0.1295 | 0.165 | 0.15 | 0.19202 | 0.3135 | 0.3295 | 0.314 | 0.2725 | 0.34 | 0.30712 | 0.35 | 0.2255 | 0.2616 | 0.1031 | 0.1865 | 0.2495 | 0.2885 | 0.3065
| Std Dev. |  | 0.020091 | 0.011063 | 0.0085557 | 0.011576 | 0.00016737 | 0.0026394 | 0.010488 | 0.0089443 | 0.0051938 | 0.0045019 | 0.0098995 | 0.004535 | 0.0018348 | 0.0057764 | 0.0053229 | 0.015055 | 0.0043551 | 0.0081774 | 0.0016693 | 0.0060222 | 0.006091 | 0.021554 | 0.02041
| Rel.Std.Dev. |  | 6.324% | 3.303% | 1.816% | 7.326% | 0.07485% | 2.043% | 6.356% | 5.963% | 2.724% | 1.44% | 3% | 1.453% | 0.6725% | 1.705% | 1.735% | 4.221% | 1.928% | 3.134% | 1.628% | 3.209% | 2.451% | 7.506% | 6.559%
| PDM3 |  | 23.09% | 29.79% | 82.48% | -38.79% | -13.36% | -49.96% | -36.07% | -41.89% | -26.14% | 21.14% | 27.85% | 20.94% | 5.704% | 31.27% | 18.86% | 38.18% | -12.51% | 1.074% | -60.28% | -27.29% | -3.723% | 11.26% | 20.56%
| Z-Score (Absolute) |  | 0.53 | 0.73 | 2.34 | 1.35 | 0.58 | 1.69 | 1.27 | 1.45 | 0.97 | 0.47 | 0.67 | 0.46 | 0 | 0.78 | 0.4 | 0.99 | 0.55 | 0.14 | 2.01 | 1 | 0.29 | 0.17 | 0.45
| Analytical results for Tl in OREAS 553b (Certified Value 0.67 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 0.71 | 0.62 | 0.71 | 0.7 | 0.74 | 0.61 | 0.66 | 0.68249 | 0.67 | 0.65 | 0.7 | 0.66 | 0.67 | 0.74779 | 0.67 | 0.69 | 0.72 | 0.364 | 0.7 | 0.6 | 0.61 | 0.58 |  | 1
| 1 | 2 | 0.73 | 0.57 | 0.72 | 0.68 | 0.71 | 0.65 | 0.62 | 0.69144 | 0.66 | 0.69 | 0.68 | 0.68 | 0.65 | 0.7029 | 0.66 | 0.67 | 0.71 | 0.31 | 0.6 | 0.7 | 0.65 | 0.55 |  | 27
| 1 | 3 | 0.71 | 0.6 | 0.68 | 0.69 | 0.73 | 0.68 | 0.65 | 0.6888 | 0.66 | 0.63 | 0.66 | 0.68 | 0.66 | 0.72506 | 0.65 | 0.68 | 0.73 | 0.323 | 0.7 | 0.8 | 0.63 | 0.57 |  | 16
| 1 | 4 | 0.72 | 0.58 | 0.71 | 0.7 | 0.72 | 0.67 | 0.63 | 0.67848 | 0.64 | 0.66 | 0.64 | 0.66 | 0.63 | 0.7473 | 0.65 | 0.72 | 0.73 | 0.566 | 0.6 | 0.7 | 0.55 | 0.57 |  | 0.66915
| 1 | 5 | 0.72 | 0.62 | 0.73 | 0.69 | 0.76 | 0.61 | 0.61 | 0.7015 | 0.64 | 0.66 | 0.67 | 0.68 | 0.64 | 0.7283 | 0.65 | 0.69 | 0.73 | 0.35 | 0.6 | 0.7 | 0.65 | 0.57 |  | 57
| 1 | 6 | 0.69 | 0.59 | 0.71 | 0.7 | 0.72 | 0.61 | 0.65 | 0.70082 | 0.7 | 0.66 | 0.68 | 0.7 | 0.64 | 0.71807 | 0.66 | 0.7 | 0.74 | 0.313 | 0.7 | 0.6 | 0.63 | 0.55
| Mean |  | 0.71333 | 0.59667 | 0.71 | 0.69333 | 0.73 | 0.63833 | 0.63667 | 0.69059 | 0.66167 | 0.65833 | 0.67167 | 0.67667 | 0.64833 | 0.72824 | 0.65667 | 0.69167 | 0.72667 | 0.371 | 0.65 | 0.68333 | 0.62 | 0.565
| Median |  | 0.715 | 0.595 | 0.71 | 0.695 | 0.725 | 0.63 | 0.64 | 0.69012 | 0.66 | 0.66 | 0.675 | 0.68 | 0.645 | 0.72668 | 0.655 | 0.69 | 0.73 | 0.3365 | 0.65 | 0.7 | 0.63 | 0.57
| Std Dev. |  | 0.013663 | 0.020656 | 0.016733 | 0.008165 | 0.017889 | 0.032506 | 0.019664 | 0.009378 | 0.022286 | 0.019408 | 0.020412 | 0.015055 | 0.01472 | 0.017327 | 0.008165 | 0.017224 | 0.010328 | 0.097882 | 0.054772 | 0.075277 | 0.037417 | 0.012247
| Rel.Std.Dev. |  | 1.915% | 3.462% | 2.357% | 1.178% | 2.45% | 5.092% | 3.089% | 1.358% | 3.368% | 2.948% | 3.039% | 2.225% | 2.27% | 2.379% | 1.243% | 2.49% | 1.421% | 26.38% | 8.427% | 11.02% | 6.035% | 2.168%
| PDM3 |  | 6.603% | -10.83% | 6.105% | 3.614% | 9.094% | -4.605% | -4.854% | 3.204% | -1.118% | -1.616% | 0.3763% | 1.123% | -3.111% | 8.831% | -1.865% | 3.365% | 8.596% | -44.56% | -2.862% | 2.12% | -7.345% | -15.56%
| Z-Score (Absolute) |  | 1.08 | 1.62 | 1 | 0.62 | 1.46 | 0.66 | 0.69 | 0.55 | 0.12 | 0.19 | 0.12 | 0.23 | 0.42 | 1.42 | 0.23 | 0.58 | 1.39 | 6.84 | NA | NA | 1.08 | 2.35
Results from laboratories 20 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 553b (Certified Value 0.42 ppm) | [24 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 15 | 25 | 26 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.41 | 0.3 | 0.5 | 0.57528 | 0.45 | 0.43032 | 0.41 | 0.29 | [15 empty] | 1
| 1 | 2 | 0.42 | 0.3 | 0.5 | 0.5464 | 0.41 | 0.41484 | 0.41 | 0.29 | [15 empty] | 12
| 1 | 3 | 0.41 | 0.3 | 0.5 | 0.51744 | 0.43 | 0.44092 | 0.41 | 0.29 | [15 empty] | 16
| 1 | 4 | 0.44 | 0.3 | 0.5 | 0.5256 | 0.41 | 0.42705 | 0.44 | 0.3 | [15 empty] | 0.4214
| 1 | 5 | 0.42 | 0.4 | 0.5 | 0.59232 | 0.43 | 0.41782 | 0.41 | 0.28 | [15 empty] | 58
| 1 | 6 | 0.44 | 0.3 | 0.5 | 0.55288 | 0.4 | 0.42957 | 0.41 | 0.26
| Mean |  | 0.42333 | 0.31667 | 0.5 | 0.55165 | 0.42167 | 0.42675 | 0.415 | 0.285
| Median |  | 0.42 | 0.3 | 0.5 | 0.54964 | 0.42 | 0.42831 | 0.41 | 0.29
| Std Dev. |  | 0.013663 | 0.040825 | 0 | 0.028602 | 0.018348 | 0.0094154 | 0.012247 | 0.013784
| Rel.Std.Dev. |  | 3.227% | 12.89% | 0% | 5.185% | 4.351% | 2.206% | 2.951% | 4.837%
| PDM3 |  | 0.4586% | -24.85% | 18.65% | 30.91% | 0.06308% | 1.27% | -1.519% | -32.37%
| Z-Score (Absolute) |  | 0.1 | NA | NA | 14.82 | 0.1 | 0.49 | 0.86 | 15.78
Results from laboratories 2 and 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 553b (Certified Value 8.6 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 8.58 | 7 | 9 | 8.27 | 10.3 | 8.1 | 8.5 | 8.7715 | 8.1 | 8.2 | 8.6 | 8.4 | 8.7 | 9.0721 | 8.56 | 8.47 | 9.26 | 9.82 | 6.7 | 8.1 | 9.12 | < 10 |  | 1
| 1 | 2 | 8.58 | 7.1 | 9.1 | 8.23 | 10.3136 | 8 | 9 | 8.3673 | 8 | 8.9 | 8.4 | 8.6 | 8.4 | 8.86 | 8.55 | 8.48 | 9.24 | 9.14 | 7.1 | 8.5 | 9.32 | < 10 |  | 13
| 1 | 3 | 8.62 | 7.3 | 8.9 | 8.17 | 10.3152 | 8.2 | 8.7 | 8.8228 | 8.3 | 8.4 | 8.7 | 8.2 | 8.7 | 9.2046 | 8.45 | 8.85 | 9 | 7.97 | 7.3 | 8.6 | 9.08 | < 10 |  | 16
| 1 | 4 | 8.53 | 6.9 | 9.1 | 8.41 | 10.3392 | 7.9 | 8.4 | 8.4364 | 8.1 | 7.5 | 8.6 | 8.2 | 8.4 | 9.1549 | 8.61 | 8.79 | 9.13 | 5.78 | 7.4 | 8.8 | 9.28 | < 10 |  | 8.6004
| 1 | 5 | 8.7 | 7.8 | 9.3 | 8.23 | 10.3824 | 7.7 | 8.9 | 8.383 | 7.8 | 8.5 | 8.7 | 8.2 | 8.7 | 9.1025 | 8.58 | 8.9 | 9.28 | 7.92 | 7.4 | 8.4 | 8.94 | < 10 |  | 59
| 1 | 6 | 8.76 | 7.2 | 9 | 8.16 | 10.324 | 7.8 | 8.6 | 8.7136 | 8.4 | 8.5 | 8.6 | 8.7 | 8.4 | 9.035 | 8.58 | 8.88 | 9.11 | 7.72 | 7.8 | 8.5 | 9.14 | < 10
| Mean |  | 8.6283 | 7.2167 | 9.0667 | 8.245 | 10.3291 | 7.95 | 8.6833 | 8.5824 | 8.1167 | 8.3333 | 8.6 | 8.3833 | 8.55 | 9.0715 | 8.555 | 8.7283 | 9.17 | 8.0583 | 7.2833 | 8.4833 | 9.1467
| Median |  | 8.6 | 7.15 | 9.05 | 8.23 | 10.3196 | 7.95 | 8.65 | 8.575 | 8.1 | 8.45 | 8.6 | 8.3 | 8.55 | 9.0873 | 8.57 | 8.82 | 9.185 | 7.945 | 7.35 | 8.5 | 9.13
| Std Dev. |  | 0.085888 | 0.31885 | 0.13663 | 0.090719 | 0.029152 | 0.18708 | 0.23166 | 0.20886 | 0.2137 | 0.46762 | 0.10954 | 0.22286 | 0.16432 | 0.11973 | 0.055408 | 0.19974 | 0.10881 | 1.3861 | 0.3656 | 0.23166 | 0.13837
| Rel.Std.Dev. |  | 0.9954% | 4.418% | 1.507% | 1.1% | 0.2822% | 2.353% | 2.668% | 2.434% | 2.633% | 5.611% | 1.274% | 2.658% | 1.922% | 1.32% | 0.6477% | 2.288% | 1.187% | 17.2% | 5.02% | 2.731% | 1.513%
| PDM3 |  | 0.3243% | -16.09% | 5.421% | -4.133% | 20.1% | -7.563% | 0.9638% | -0.2094% | -5.625% | -3.106% | -0.005142% | -2.524% | -0.5865% | 5.477% | -0.5284% | 1.487% | 6.622% | -6.303% | -15.31% | -1.362% | 6.351%
| Z-Score (Absolute) |  | 0.14 | 2.42 | 0.93 | 0.56 | 3.22 | 1.09 | 0.24 | 0.05 | 0.79 | 0.4 | 0.09 | 0.31 | 0 | 0.94 | 0 | 0.32 | 1.12 | 0.9 | 2.3 | 0.13 | 1.08 | 6.45
| Analytical results for V in OREAS 553b (Certified Value 348 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 351 | 347 | 398 | 351 | 333.65 | 364 | 339 | 325 | 364.799 | 343 | 349 | 349 | 342 | 345 | 354.6127 | 362 | 343 | 369 | 356 | 328 | 349 | 348 | 376 | 1
| 1 | 2 | 346 | 341 | 401 | 348 | 333.55 | 380 | 340 | 335 | 355.2428 | 339 | 350 | 338 | 340 | 340 | 344.951 | 367 | 341 | 373 | 353 | 330 | 355 | 348 | 372 | 29
| 1 | 3 | 350 | 345 | 391 | 344 | 333.13 | 357 | 340 | 320 | 363.224 | 340 | 336 | 341 | 340 | 337 | 359.4149 | 356 | 337 | 372 | 353 | 328 | 352 | 349 | 372 | 16
| 1 | 4 | 351 | 346 | 402 | 344 | 329.74 | 362 | 339 | 334 | 349.0287 | 340 | 341 | 347 | 340 | 340 | 353.6073 | 368 | 343 | 371 | 343 | 324 | 359 | 383 | 382 | 348.4629
| 1 | 5 | 353 | 359 | 400 | 349 | 338.57 | 346 | 336 | 326 | 350.1358 | 341 | 341 | 340 | 349 | 333 | 355.3002 | 364 | 347 | 372 | 346 | 321 | 357 | 337 | 380 | 60
| 1 | 6 | 361 | 343 | 390 | 343 | 335.48 | 339 | 337 | 321 | 355.9081 | 345 | 351 | 345 | 358 | 335 | 355.7592 | 364 | 346 | 373 | 361 | 329 | 362 | 346 | 377
| Mean |  | 352 | 346.8333 | 397 | 346.5 | 334.02 | 358 | 338.5 | 326.8333 | 356.3897 | 341.3333 | 344.6667 | 343.3333 | 344.8333 | 338.3333 | 353.9409 | 363.5 | 342.8333 | 371.6667 | 352 | 326.6667 | 355.6667 | 351.8333 | 376.5
| Median |  | 351 | 345.5 | 399 | 346 | 333.6 | 359.5 | 339 | 325.5 | 355.5755 | 340.5 | 345 | 343 | 341 | 338.5 | 354.9565 | 364 | 343 | 372 | 353 | 328 | 356 | 348 | 376.5
| Std Dev. |  | 4.98 | 6.3377 | 5.2154 | 3.2711 | 2.9087 | 14.4361 | 1.6432 | 6.3692 | 6.5156 | 2.2509 | 6.1536 | 4.3205 | 7.3326 | 4.274 | 4.827 | 4.2778 | 3.6009 | 1.5055 | 6.5727 | 3.4448 | 4.7188 | 15.8923 | 4.0866
| Rel.Std.Dev. |  | 1.415% | 1.827% | 1.314% | 0.944% | 0.8708% | 4.032% | 0.4854% | 1.949% | 1.828% | 0.6595% | 1.785% | 1.258% | 2.126% | 1.263% | 1.364% | 1.177% | 1.05% | 0.4051% | 1.867% | 1.055% | 1.327% | 4.517% | 1.085%
| PDM3 |  | 1.015% | -0.4676% | 13.93% | -0.5633% | -4.145% | 2.737% | -2.859% | -6.207% | 2.275% | -2.046% | -1.089% | -1.472% | -1.042% | -2.907% | 1.572% | 4.315% | -1.616% | 6.659% | 1.015% | -6.255% | 2.067% | 0.9672% | 8.046%
| Z-Score (Absolute) |  | 0.42 | 0 | 4.06 | 0.03 | 1.04 | 0.9 | 0.67 | 1.62 | 0.77 | 0.45 | 0.18 | 0.28 | 0.16 | 0.69 | 0.58 | 1.35 | 0.32 | 2.01 | 0.42 | 1.63 | 0.71 | 0.4 | 2.4
| Analytical results for W in OREAS 553b (Certified Value 2.52 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.4 | 2.7 | 3.7 | 2 | 2.6 | 2.3 | 1.9155 | 2.7 | 2.6 | 2.7 | 2.6 | 2.6 | 2.766 | 2.68 | 2.3 | 2.5 | 3.89 | 1.9 | 2 | 2.8 | 2.4 |  |  | 1
| 1 | 2 | 2.9 | 2.5 | 4.4 | 2.1 | 2.6 | 2.1 | 1.892 | 2.6 | 2.8 | 2.7 | 2.6 | 2.5 | 2.6908 | 2.53 | 2.3 | 2.5 | 4.21 | 2 | 2.2 | 2.9 | 2.3 |  |  | 14
| 1 | 3 | 2.6 | 2.5 | 8.8 | 2 | 2.3 | 1.9 | 1.8692 | 2.6 | 2.5 | 2.7 | 2.5 | 2.5 | 2.9045 | 2.69 | 2.5 | 2.6 | 4 | 2 | 2.1 | 2.7 | 2.4 |  |  | 16
| 1 | 4 | 2.7 | 2.6 | 3.7 | 2 | 2.4 | 2.2 | 1.8209 | 2.7 | 2.6 | 2.8 | 2.5 | 2.5 | 2.8659 | 2.55 | 2.3 | 2.5 | 3.63 | 2 | 2.1 | 2.5 | 2.5 |  |  | 2.52
| 1 | 5 | 2.6 | 2.5 | 3.7 | 2.1 | 2.3 | 2 | 1.8692 | 2.6 | 2.6 | 2.8 | 2.5 | 2.5 | 2.762 | 2.68 | 2.4 | 2.5 | 4.83 | 2 | 2.1 | 2.3 | 2.5 |  |  | 61
| 1 | 6 | 2.7 | 2.4 | 3.5 | 2 | 2.4 | 2.2 | 1.8037 | 2.8 | 2.6 | 2.7 | 2.5 | 2.5 | 2.7687 | 2.63 | 2.4 | 2.6 | 4.5 | 2.1 | 2.1 | 2.8 | 2.5
| Mean |  | 2.65 | 2.5333 | 4.6333 | 2.0333 | 2.4333 | 2.1167 | 1.8617 | 2.6667 | 2.6167 | 2.7333 | 2.5333 | 2.5167 | 2.793 | 2.6267 | 2.3667 | 2.5333 | 4.1767 | 2 | 2.1 | 2.6667 | 2.4333
| Median |  | 2.65 | 2.5 | 3.7 | 2 | 2.4 | 2.15 | 1.8692 | 2.65 | 2.6 | 2.7 | 2.5 | 2.5 | 2.7673 | 2.655 | 2.35 | 2.5 | 4.105 | 2 | 2.1 | 2.75 | 2.45
| Std Dev. |  | 0.16432 | 0.10328 | 2.0646 | 0.05164 | 0.13663 | 0.1472 | 0.042321 | 0.08165 | 0.098319 | 0.05164 | 0.05164 | 0.040825 | 0.078064 | 0.070616 | 0.08165 | 0.05164 | 0.43468 | 0.063246 | 0.063246 | 0.22509 | 0.08165
| Rel.Std.Dev. |  | 6.201% | 4.077% | 44.56% | 2.54% | 5.615% | 6.954% | 2.273% | 3.062% | 3.757% | 1.889% | 2.038% | 1.622% | 2.795% | 2.688% | 3.45% | 2.038% | 10.41% | 3.162% | 3.012% | 8.441% | 3.355%
| PDM3 |  | 5.16% | 0.53% | 83.86% | -19.31% | -3.438% | -16% | -26.12% | 5.821% | 3.837% | 8.467% | 0.53% | -0.1314% | 10.83% | 4.234% | -6.084% | 0.53% | 65.74% | -20.63% | -16.67% | 5.821% | -3.438%
| Z-Score (Absolute) |  | 0.59 | 0 | 10.62 | 2.53 | 0.51 | 2.11 | 3.4 | 0.67 | 0.42 | 1.01 | 0 | 0.08 | 1.31 | 0.47 | 0.84 | 0 | 8.31 | 2.7 | 2.19 | 0.67 | 0.51
| Analytical results for Y in OREAS 553b (Certified Value 23.8 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 22.92 | 21.8 | 27.5 | 23.2 | 32.4427 | 25.5 | 24 | 21.2 | 22.7869 | 23.4 | 23.2 | 24 | 25.6 | 24.9 | 24.0382 | 23.4 | 22.5 | 26.25 | 22.19 | 18.3 | 22.5 | 23.8 | 20 | 1
| 1 | 2 | 22.85 | 21.6 | 28.8 | 23.1 | 32.3853 | 22.7 | 24 | 20.4 | 23.9334 | 22.3 | 25.9 | 24.1 | 25.7 | 23.7 | 24.4417 | 23.9 | 23.4 | 25.9 | 21.98 | 18.8 | 24 | 23.6 | 19.8 | 15
| 1 | 3 | 23.24 | 21.5 | 27.9 | 23 | 32.4025 | 23.9 | 24 | 21.7 | 23.1168 | 24.4 | 22.7 | 23.6 | 25.2 | 23.8 | 24.8272 | 23.5 | 22.9 | 25.7 | 22.26 | 18.2 | 24.4 | 23.5 | 19.4 | 16
| 1 | 4 | 24.53 | 21.4 | 28.6 | 23.4 | 32.3942 | 23.5 | 24 | 20.5 | 24.5154 | 22.4 | 21.6 | 23.7 | 25.7 | 24.3 | 24.4276 | 24.1 | 24 | 25.96 | 22.07 | 19.3 | 26 | 25.8 | 19.7 | 23.7694
| 1 | 5 | 23.45 | 23.3 | 28.9 | 23.3 | 32.3817 | 24.7 | 24 | 21.8 | 23.8367 | 21.9 | 23.8 | 23.5 | 25.1 | 24 | 24.2992 | 24.2 | 23.6 | 25.99 | 21.9 | 21.1 | 23.9 | 23 | 18.8 | 62
| 1 | 6 | 23.75 | 22.2 | 28.2 | 23.8 | 32.428 | 22.8 | 24 | 20.3 | 24.6679 | 25.3 | 23.8 | 23.7 | 25.9 | 23.7 | 23.9324 | 24.3 | 23.7 | 25.68 | 23 | 19.8 | 24 | 23.1 | 17.4
| Mean |  | 23.4567 | 21.9667 | 28.3167 | 23.3 | 32.4057 | 23.85 | 24 | 20.9833 | 23.8095 | 23.2833 | 23.5 | 23.7667 | 25.5333 | 24.0667 | 24.3277 | 23.9 | 23.35 | 25.9133 | 22.2333 | 19.25 | 24.1333 | 23.8 | 19.1833
| Median |  | 23.345 | 21.7 | 28.4 | 23.25 | 32.3983 | 23.7 | 24 | 20.85 | 23.8851 | 22.9 | 23.5 | 23.7 | 25.65 | 23.9 | 24.3634 | 24 | 23.5 | 25.93 | 22.13 | 19.05 | 24 | 23.55 | 19.55
| Std Dev. |  | 0.62308 | 0.71181 | 0.54924 | 0.28284 | 0.024506 | 1.095 | 0 | 0.67355 | 0.74523 | 1.3408 | 1.4339 | 0.23381 | 0.31411 | 0.46762 | 0.32041 | 0.37417 | 0.55408 | 0.21049 | 0.39808 | 1.0895 | 1.1237 | 1.0257 | 0.96833
| Rel.Std.Dev. |  | 2.656% | 3.24% | 1.94% | 1.214% | 0.07562% | 4.591% | 0% | 3.21% | 3.13% | 5.759% | 6.102% | 0.9838% | 1.23% | 1.943% | 1.317% | 1.566% | 2.373% | 0.8123% | 1.79% | 5.66% | 4.656% | 4.31% | 5.048%
| PDM3 |  | -1.316% | -7.584% | 19.13% | -1.975% | 36.33% | 0.339% | 0.9701% | -11.72% | 0.1687% | -2.045% | -1.133% | -0.01154% | 7.421% | 1.251% | 2.349% | 0.5494% | -1.764% | 9.02% | -6.462% | -19.01% | 1.531% | 0.1287% | -19.29%
| Z-Score (Absolute) |  | 0.46 | 2.47 | 6.09 | 0.67 | 11.61 | 0.07 | 0.27 | 3.8 | 0.01 | 0.7 | 0.4 | 0.04 | 2.34 | 0.36 | 0.71 | 0.13 | 0.61 | 2.85 | 2.11 | 6.14 | 0.45 | 0 | 6.23
| Analytical results for Yb in OREAS 553b (Certified Value 2.84 ppm) | [24 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 15 | 16 | 17 | 25 | 26 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
2
| 1 | 1 | 2.76 | 2.5 | 3.2 | 3.5659 | 3.3 | 2.7 | 3.0092 | 2.69 | 2.7 | 2.64 | 2.1 | [12 empty] | 1
| 1 | 2 | 2.81 | 2.4 | 3.3 | 3.5198 | 3.3 | 2.3 | 2.8462 | 2.77 | 2.8 | 2.65 | 2.1 | [12 empty] | 16
| 1 | 3 | 2.77 | 2.5 | 3.2 | 3.5406 | 3.2 | 2.6 | 3.0847 | 2.84 | 2.7 | 2.7 | 2.1 | [12 empty] | 16
| 1 | 4 | 2.84 | 2.5 | 3.3 | 3.562 | 3.2 | 2.3 | 3.0345 | 2.84 | 2.8 | 2.82 | 2.2 | [12 empty] | 2.8396
| 1 | 5 | 2.94 | 2.6 | 3.3 | 3.591 | 3.3 | 2.6 | 2.9071 | 2.8 | 2.9 | 2.63 | 2.1 | [12 empty] | 63
| 1 | 6 | 2.91 | 2.5 | 3.3 | 3.5824 | 3.2 | 2.4 | 2.9976 | 2.81 | 2.8 | 2.75 | 1.9
| Mean |  | 2.8383 | 2.5 | 3.2667 | 3.5603 | 3.25 | 2.4833 | 2.9799 | 2.7917 | 2.7833 | 2.6983 | 2.0833
| Median |  | 2.825 | 2.5 | 3.3 | 3.564 | 3.25 | 2.5 | 3.0034 | 2.805 | 2.8 | 2.675 | 2.1
| Std Dev. |  | 0.073598 | 0.063246 | 0.05164 | 0.026443 | 0.054772 | 0.17224 | 0.087522 | 0.056362 | 0.075277 | 0.074677 | 0.098319
| Rel.Std.Dev. |  | 2.593% | 2.53% | 1.581% | 0.7427% | 1.685% | 6.936% | 2.937% | 2.019% | 2.705% | 2.768% | 4.719%
| PDM3 |  | -0.04315% | -11.96% | 15.04% | 25.38% | 14.45% | -12.55% | 4.942% | -1.687% | -1.98% | -4.973% | -26.63%
| Z-Score (Absolute) |  | 0.11 | 0.67 | 1.1 | 1.78 | 1.06 | 0.71 | 0.44 | 0 | 0.02 | 0.22 | 1.64
| Analytical results for Zn in OREAS 553b (Certified Value 16.1 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 17 | 37.6 | 16 | 17 | 18 | 12 | 15 | 15.1227 | 15 | 19 | 16 | 15 | 16 | 16.3875 | 16.5 | 16 | 19 | 16 | 15 | 18 | 17 | 14 |  | 1
| 1 | 2 | 16 | 34.6 | 17 | 16 | 20.76 | 12 | 14 | 15.5917 | 15 | 20 | 15 | 16 | 15 | 16.2577 | 16.7 | 16 | 18 | 16 | 15 | 16 | 16.3 | 15 |  | 17
| 1 | 3 | 17 | 36.1 | 16 | 17 | 18.18 | 11 | 15 | 15.1174 | 15 | 19 | 15 | 16 | 16 | 16.802 | 16.6 | 16 | 18 | 16 | 15 | 17 | 15.4 | 17 |  | 16
| 1 | 4 | 16 | 35.3 | 16 | 17 | 20.1 | 12 | 14 | 15.0873 | 15 | 19 | 15 | 16 | 16 | 16.3662 | 16.6 | 16 | 18 | 15 | 15 | 19 | 16.8 | 18 |  | 16.1493
| 1 | 5 | 17 | 36.7 | 16 | 17 | 21.1 | 11 | 14 | 15.7264 | 14 | 17 | 15 | 18 | 16 | 16.2466 | 16.6 | 16 | 17 | 15 | 14 | 18 | 14.4 | 15 |  | 64
| 1 | 6 | 16 | 34.6 | 16 | 17 | 18.98 | 11 | 15 | 14.7226 | 15 | 19 | 15 | 17 | 16 | 16.2951 | 16.8 | 16 | 18 | 16 | 15 | 18 | 15 | 16
| Mean |  | 16.5 | 35.8167 | 16.1667 | 16.8333 | 19.52 | 11.5 | 14.5 | 15.228 | 14.8333 | 18.8333 | 15.1667 | 16.3333 | 15.8333 | 16.3925 | 16.6333 | 16 | 18 | 15.6667 | 14.8333 | 17.6667 | 15.8167 | 15.8333
| Median |  | 16.5 | 35.7 | 16 | 17 | 19.54 | 11.5 | 14.5 | 15.12 | 15 | 19 | 15 | 16 | 16 | 16.3307 | 16.6 | 16 | 18 | 16 | 15 | 18 | 15.85 | 15.5
| Std Dev. |  | 0.54772 | 1.2057 | 0.40825 | 0.40825 | 1.3242 | 0.54772 | 0.54772 | 0.36858 | 0.40825 | 0.98319 | 0.40825 | 1.0328 | 0.40825 | 0.20848 | 0.10328 | 0 | 0.63246 | 0.5164 | 0.40825 | 1.0328 | 1.0439 | 1.472
| Rel.Std.Dev. |  | 3.32% | 3.366% | 2.525% | 2.425% | 6.784% | 4.763% | 3.777% | 2.42% | 2.752% | 5.22% | 2.692% | 6.323% | 2.578% | 1.272% | 0.6209% | 0% | 3.514% | 3.296% | 2.752% | 5.846% | 6.6% | 9.297%
| PDM3 |  | 2.171% | 121.8% | 0.1074% | 4.236% | 20.87% | -28.79% | -10.21% | -5.705% | -8.149% | 16.62% | -6.085% | 1.139% | -1.957% | 1.506% | 2.997% | -0.9247% | 11.46% | -2.989% | -8.149% | 9.396% | -2.06% | -1.957%
| Z-Score (Absolute) |  | 0.35 | 16.58 | 0.07 | 0.63 | 2.89 | 3.85 | 1.33 | 0.72 | 1.05 | 2.31 | 0.77 | 0.21 | 0.21 | 0.26 | 0.46 | 0.07 | 1.61 | 0.35 | 1.05 | 1.33 | 0.22 | 0.21
| Analytical results for Zr in OREAS 553b (Certified Value 188 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 0
0
| 1 | 1 | 186 | 181.4 | 197 | 193 | 170.29 | 181 | 200 | 194.2 | 173.8671 | 198 | 195.5 | 182.5 | 192.5 | 198.5 | 190.2703 | 195 | 190 | 192.3 | 174.9 | 172.9 | 179 | 202 | 188 | 1
| 1 | 2 | 185.6 | 172.9 | 198 | 197 | 173.32 | 190 | 192 | 186.1 | 178.7965 | 194 | 204 | 181.5 | 189.5 | 193.5 | 190.8282 | 195 | 192 | 199.3 | 173.1 | 174.1 | 175.7 | 198 | 191 | 21
| 1 | 3 | 189.9 | 175.8 | 190 | 192 | 173.1 | 180 | 201 | 192.5 | 177.2341 | 197 | 193 | 178.5 | 190 | 195 | 194.2362 | 197 | 189 | 190.3 | 177.3 | 170.9 | 176.2 | 201 | 190 | 16
| 1 | 4 | 193.6 | 179.4 | 196 | 193 | 170.8 | 180 | 193 | 198 | 170.6553 | 194.5 | 194.5 | 184.5 | 192.5 | 193 | 192.6722 | 199 | 193 | 192.8 | 170.9 | 169.7 | 177 | 216 | 191 | 187.8487
| 1 | 5 | 189.1 | 184.2 | 201 | 197 | 171.47 | 175 | 191 | 193.1 | 166.242 | 191.5 | 197.5 | 182 | 195 | 193 | 189.9312 | 197 | 189 | 193.2 | 174.6 | 172 | 175.7 | 200 | 190 | 65
| 1 | 6 | 194.6 | 175.9 | 197 | 196 | 169.86 | 171 | 199 | 185.4 | 169.7604 | 202 | 199.5 | 182 | 194 | 193 | 188.8929 | 199 | 192 | 192.1 | 177 | 173.4 | 175.3 | 201 | 187
| Mean |  | 189.8 | 178.2667 | 196.5 | 194.6667 | 171.4733 | 179.5 | 196 | 191.55 | 172.7592 | 196.1667 | 197.3333 | 181.8333 | 192.25 | 194.3333 | 191.1385 | 197 | 190.8333 | 193.3333 | 174.6333 | 172.1667 | 176.4833 | 203 | 189.5
| Median |  | 189.5 | 177.65 | 197 | 194.5 | 171.135 | 180 | 196 | 192.8 | 172.2612 | 195.75 | 196.5 | 182 | 192.5 | 193.25 | 190.5492 | 197 | 191 | 192.55 | 174.75 | 172.45 | 175.95 | 201 | 190
| Std Dev. |  | 3.7433 | 4.1664 | 3.6194 | 2.2509 | 1.45 | 6.4109 | 4.4721 | 4.8878 | 4.7664 | 3.6697 | 3.9833 | 1.9408 | 2.1622 | 2.1833 | 1.9646 | 1.7889 | 1.7224 | 3.0885 | 2.4114 | 1.644 | 1.3644 | 6.5115 | 1.6432
| Rel.Std.Dev. |  | 1.972% | 2.337% | 1.842% | 1.156% | 0.8456% | 3.572% | 2.282% | 2.552% | 2.759% | 1.871% | 2.019% | 1.067% | 1.125% | 1.123% | 1.028% | 0.908% | 0.9026% | 1.597% | 1.381% | 0.9549% | 0.7731% | 3.208% | 0.8671%
| PDM3 |  | 1.039% | -5.101% | 4.605% | 3.629% | -8.717% | -4.444% | 4.339% | 1.97% | -8.033% | 4.428% | 5.049% | -3.202% | 2.343% | 3.452% | 1.751% | 4.872% | 1.589% | 2.92% | -7.035% | -8.348% | -6.05% | 8.066% | 0.879%
| Z-Score (Absolute) |  | 0.17 | 1.62 | 0.67 | 0.44 | 2.47 | 1.46 | 0.61 | 0.05 | 2.31 | 0.63 | 0.78 | 1.17 | 0.14 | 0.4 | 0 | 0.74 | 0.04 | 0.28 | 2.08 | 2.39 | 1.84 | 1.49 | 0.21
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C634: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C652: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C670: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C688: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C707: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C725: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C743: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C761: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C870: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C906: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C943: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C961: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach (05%)
| Analytical results for Co in OREAS 553b (Certified Value 0.127 wt.%) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 09 | 19 | 23 | 25 | [11 empty] | wt.%
|  |  | AL-5S*OES | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | [11 empty] | 3
3
| 1 | 1 | 0.1246 | 0.139 | 0.18895 | 0.12317 | 0.1229 | 0.13477 | 0.13 | 0.11 | 2.501 | [11 empty] | 1
| 1 | 2 | 0.1247 | 0.139 | 0.19143 | 0.12115 | 0.1243 | 0.13551 | 0.132 | 0.12 | 2.604 | [11 empty] | 30
| 1 | 3 | 0.1242 | 0.14 | 0.19181 | 0.12209 | 0.1257 | 0.13243 | 0.132 | 0.11 | 2.589 | [11 empty] | 16
| 1 | 4 | 0.1232 | 0.138 | 0.18699 | 0.12269 | 0.1262 | 0.13449 | 0.132 | 0.12 | 2.665 | [11 empty] | 0.12698
| 1 | 5 | 0.1253 | 0.14 | 0.1878 | 0.12108 | 0.1236 | 0.13326 | 0.129 | 0.1 | 2.596 | [11 empty] | 67
| 1 | 6 | 0.1243 | 0.14 | 0.19056 | 0.1233 | 0.1246 | 0.13344 | 0.127 | 0.11 | 2.545
| Mean |  | 0.12438 | 0.13933 | 0.18959 | 0.12225 | 0.12455 | 0.13398 | 0.13033 | 0.11167 | 2.5833
| Median |  | 0.12445 | 0.1395 | 0.18975 | 0.12239 | 0.12445 | 0.13396 | 0.131 | 0.11 | 2.5925
| Std Dev. |  | 0.0006969 | 0.0008165 | 0.0019828 | 0.00097657 | 0.0012438 | 0.0011348 | 0.0020656 | 0.0075277 | 0.055745
| Rel.Std.Dev. |  | 0.5603% | 0.586% | 1.046% | 0.7989% | 0.9986% | 0.847% | 1.585% | 6.741% | 2.158%
| PDM3 |  | -2.041% | 9.733% | 49.31% | -3.725% | -1.91% | 5.518% | 2.645% | -12.06% | 1935%
| Z-Score (Absolute) |  | 0.5 | 0.75 | 4.94 | 0.67 | 0.48 | 0.3 | 0 | 1.56 | 204.51
| Analytical results for Cu in OREAS 553b (Certified Value 1.07 wt.%) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 19 | 20 | 23 | 25 | 26 | wt.%
|  |  | AL-5S*OES | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | 2
3
| 1 | 1 | 1.0823 | 1.018 | 1.12 | 1.07 | 2.0184 | 1.0012 | 1.19 | 1.078 | 0.99741 | 0.951 | 1.01 | 1.005 | 1.115 | 0.97668 | 1.138 | 1.012 | 1.062 | 1.27 | 15.77 | 1.062 | 1
| 1 | 2 | 1.0852 | 1.038 | 1.11 | 1.07 | 2.0359 | 1.0013 | 1.188 | 1.094 | 0.97267 | 0.969 | 1.015 | 0.977 | 1.16 | 0.97031 | 1.146 | 1.036 | 1.085 | 1.24 | 16 | 1.047 | 31
| 1 | 3 | 1.0758 | 1.032 | 1.12 | 1.08 | 2.0568 | 1.0438 | 1.1966 | 1.085 | 0.98436 | 0.944 | 1.04 | 0.976 | 1.15 | 0.97379 | 1.148 | 1.037 | 1.061 | 1.26 | 16.42 | 1.06 | 16
| 1 | 4 | 1.0692 | 1.014 | 1.12 | 1.05 | 2.0752 | 1.0487 | 1.1726 | 1.079 | 0.99593 | 0.947 | 1.015 | 1.005 | 1.12 | 0.97893 | 1.119 | 1.045 | 1.075 | 1.25 | 16.36 | 1.07 | 1.0662
| 1 | 5 | 1.0838 | 1.002 | 1.12 | 1.07 | 2.0295 | 1.0222 | 1.1844 | 1.092 | 0.98858 | 0.965 | 1 | 0.962 | 1.175 | 0.96471 | 1.12 | 1.018 | 1.082 | 1.26 | 15.91 | 1.076 | 68
| 1 | 6 | 1.0642 | 1.024 | 1.13 | 1.07 | 2.049 | 1.0301 | 1.1814 | 1.089 | 0.97576 | 0.943 | 1.05 | 0.969 | 1.115 | 0.96573 | 1.141 | 0.998 | 1.081 | 1.27 | 16.01 | 1.066
| Mean |  | 1.0767 | 1.0213 | 1.12 | 1.0683 | 2.0441 | 1.0246 | 1.1855 | 1.0862 | 0.98579 | 0.95317 | 1.0217 | 0.98233 | 1.1392 | 0.97169 | 1.1353 | 1.0243 | 1.0743 | 1.2583 | 16.0783 | 1.0635
| Median |  | 1.0791 | 1.021 | 1.12 | 1.07 | 2.0424 | 1.0262 | 1.1862 | 1.087 | 0.98647 | 0.949 | 1.015 | 0.9765 | 1.135 | 0.97205 | 1.1395 | 1.027 | 1.078 | 1.26 | 16.005 | 1.064
| Std Dev. |  | 0.0085711 | 0.012941 | 0.0063246 | 0.0098319 | 0.020475 | 0.020386 | 0.0081748 | 0.0066758 | 0.010203 | 0.011143 | 0.019149 | 0.01837 | 0.025965 | 0.0057936 | 0.01277 | 0.017941 | 0.010463 | 0.01169 | 0.25701 | 0.0099146
| Rel.Std.Dev. |  | 0.796% | 1.267% | 0.5647% | 0.9203% | 1.002% | 1.99% | 0.6896% | 0.6146% | 1.035% | 1.169% | 1.874% | 1.87% | 2.279% | 0.5962% | 1.125% | 1.751% | 0.9739% | 0.929% | 1.599% | 0.9323%
| PDM3 |  | 0.9859% | -4.212% | 5.042% | 0.1965% | 91.71% | -3.909% | 11.19% | 1.869% | -7.545% | -10.6% | -4.18% | -7.869% | 6.84% | -8.867% | 6.48% | -3.93% | 0.7592% | 18.02% | 1408% | -0.2568%
| Z-Score (Absolute) |  | 0.06 | 0.59 | 0.58 | 0.04 | 11.51 | 0.55 | 1.35 | 0.18 | 1.01 | 1.4 | 0.59 | 1.05 | 0.8 | 1.18 | 0.76 | 0.56 | 0.04 | 2.21 | 177.53 | 0.09
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 553b (Certified Value 4.51 wt.%) | [24 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 4.52 | 4.45 | 4.56 | 4.73 | 4.77 | 4.64 | 4.54 | 3.8533 | 4.48 | 4.58 | 4.6 | 4.24 | 4.49 | 4.3586 | 4.492 | 4.04 | 4.3 | 4.52 | 4.43 | 4.28 | 5.17 | 4.63 | 4.57 | 1
| 1 | 2 | 4.53 | 4.46 | 4.55 | 4.72 | 4.8 | 4.66 | 4.52 | 3.8687 | 4.47 | 4.59 | 4.49 | 4.24 | 4.5 | 4.3656 | 4.469 | 4.13 | 4.3 | 4.56 | 4.42 | 4.24 | 5.19 | 4.62 | 4.59 | #N/A
| 1 | 3 | 4.52 | 4.43 | 4.61 | 4.74 | 4.72 | 4.61 | 4.56 | 3.8519 | 4.44 | 4.61 | 4.5 | 4.3 | 4.51 | 4.3617 | 4.464 | 4.08 | 4.28 | 4.56 | 4.41 | 4.23 | 5 | 4.62 | 4.59 | 16
| 1 | 4 | 4.54 | 4.43 | 4.54 | 4.71 | 4.77 | 4.7 | 4.51 | 3.8472 | 4.42 | 4.63 | 4.45 | 4.26 | 4.5 | 4.3606 | 4.574 | 4.07 | 4.31 | 4.56 | 4.5 | 4.24 | 5.05 | 4.61 | 4.55 | 4.5091
| 1 | 5 | 4.5 | 4.46 | 4.55 | 4.74 | 4.68 | 4.68 | 4.55 | 3.8729 | 4.44 | 4.56 | 4.53 | 4.3 | 4.48 | 4.3712 | 4.474 | 4.04 | 4.31 | 4.54 | 4.47 | 4.23 | 5.19 | 4.63 | 4.56 | 70
| 1 | 6 | 4.52 | 4.46 | 4.65 | 4.71 | 4.78 | 4.68 | 4.52 | 3.8415 | 4.47 | 4.63 | 4.47 | 4.61 | 4.52 | 4.3507 | 4.479 | 4.07 | 4.22 | 4.54 | 4.43 | 4.24 | 4.96 | 4.65 | 4.6
| Mean |  | 4.5217 | 4.4483 | 4.5767 | 4.725 | 4.7533 | 4.6617 | 4.5333 | 3.8559 | 4.4533 | 4.6 | 4.5067 | 4.325 | 4.5 | 4.3614 | 4.492 | 4.0717 | 4.2867 | 4.5467 | 4.4433 | 4.2433 | 5.0933 | 4.6267 | 4.5767
| Median |  | 4.52 | 4.455 | 4.555 | 4.725 | 4.77 | 4.67 | 4.53 | 3.8526 | 4.455 | 4.6 | 4.495 | 4.28 | 4.5 | 4.3612 | 4.4765 | 4.07 | 4.3 | 4.55 | 4.43 | 4.24 | 5.11 | 4.625 | 4.58
| Std Dev. |  | 0.013292 | 0.01472 | 0.043665 | 0.013784 | 0.044572 | 0.032506 | 0.019664 | 0.012332 | 0.023381 | 0.028284 | 0.053166 | 0.14223 | 0.014142 | 0.0068852 | 0.041304 | 0.033116 | 0.034448 | 0.01633 | 0.034448 | 0.018619 | 0.10289 | 0.013663 | 0.019664
| Rel.Std.Dev. |  | 0.294% | 0.3309% | 0.9541% | 0.2917% | 0.9377% | 0.6973% | 0.4338% | 0.3198% | 0.525% | 0.6149% | 1.18% | 3.289% | 0.3143% | 0.1579% | 0.9195% | 0.8133% | 0.8036% | 0.3592% | 0.7753% | 0.4388% | 2.02% | 0.2953% | 0.4297%
| PDM3 |  | 0.279% | -1.347% | 1.499% | 4.788% | 5.417% | 3.384% | 0.5377% | -14.49% | -1.236% | 2.016% | -0.05368% | -4.083% | -0.2015% | -3.275% | -0.3789% | -9.701% | -4.933% | 0.8334% | -1.458% | -5.894% | 12.96% | 2.608% | 1.499%
| Z-Score (Absolute) |  | 0.11 | 0.42 | 0.51 | 1.58 | 1.78 | 1.12 | 0.19 | 4.7 | 0.39 | 0.67 | 0 | 1.31 | 0.05 | 1.05 | 0.11 | 3.14 | 1.59 | 0.29 | 0.46 | 1.9 | 4.24 | 0.87 | 0.51
| Analytical results for S in OREAS 553b (Certified Value 3.4 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 3.43 | 3.33 | 3.32 | 3.46 | 3.47 | 3.39 | 3.45 | 3.3997 | 3.39 | 3.44 | 3.47 | 3.4 | 3.52 | 3.3102 | 3.344 | 3.38 | 3.52 | 3.3 | 3.48 | 3.27 | 3.3 | 3.41 | 3.29 | 1
| 1 | 2 | 3.45 | 3.38 | 3.38 | 3.46 | 3.37 | 3.4 | 3.43 | 3.4155 | 3.44 | 3.41 | 3.52 | 3.35 | 3.49 | 3.3181 | 3.296 | 3.42 | 3.54 | 3.31 | 3.47 | 3.28 | 3.4 | 3.4 | 3.32 | 31
| 1 | 3 | 3.46 | 3.32 | 3.4 | 3.49 | 2.93 | 3.4 | 3.42 | 3.4275 | 3.45 | 3.46 | 3.51 | 3.39 | 3.53 | 3.3086 | 3.294 | 3.5 | 3.59 | 3.31 | 3.47 | 3.26 | 3.2 | 3.39 | 3.32 | 16
| 1 | 4 | 3.47 | 3.37 | 3.32 | 3.46 | 3.5 | 3.39 | 3.44 | 3.4167 | 3.47 | 3.46 | 3.53 | 3.44 | 3.54 | 3.3239 | 3.279 | 3.4 | 3.58 | 3.31 | 3.49 | 3.28 | 3.3 | 3.37 | 3.27 | 3.402
| 1 | 5 | 3.49 | 3.38 | 3.35 | 3.45 | 3.2 | 3.39 | 3.41 | 3.42 | 3.44 | 3.41 | 3.49 | 3.41 | 3.53 | 3.3273 | 3.327 | 3.4 | 3.57 | 3.31 | 3.53 | 3.27 | 3.3 | 3.37 | 3.32 | 71
| 1 | 6 | 3.46 | 3.39 | 3.35 | 3.47 | 3.48 | 3.41 | 3.45 | 3.3975 | 3.38 | 3.51 | 3.52 | 3.4 | 3.52 | 3.3189 | 3.3 | 3.39 | 3.53 | 3.32 | 3.52 | 3.27 | 3.2 | 3.42 | 3.34
| Mean |  | 3.46 | 3.3617 | 3.3533 | 3.465 | 3.325 | 3.3967 | 3.4333 | 3.4128 | 3.4283 | 3.4483 | 3.5067 | 3.3983 | 3.5217 | 3.3178 | 3.3067 | 3.415 | 3.555 | 3.31 | 3.4933 | 3.2717 | 3.2833 | 3.3933 | 3.31
| Median |  | 3.46 | 3.375 | 3.35 | 3.46 | 3.42 | 3.395 | 3.435 | 3.4161 | 3.44 | 3.45 | 3.515 | 3.4 | 3.525 | 3.3185 | 3.298 | 3.4 | 3.555 | 3.31 | 3.485 | 3.27 | 3.3 | 3.395 | 3.32
| Std Dev. |  | 0.02 | 0.029269 | 0.032042 | 0.013784 | 0.22332 | 0.008165 | 0.01633 | 0.011813 | 0.035449 | 0.037639 | 0.022509 | 0.029269 | 0.017224 | 0.0073513 | 0.024047 | 0.043704 | 0.02881 | 0.0063246 | 0.02582 | 0.0075277 | 0.075277 | 0.020656 | 0.025298
| Rel.Std.Dev. |  | 0.578% | 0.8707% | 0.9555% | 0.3978% | 6.716% | 0.2404% | 0.4756% | 0.3461% | 1.034% | 1.092% | 0.6419% | 0.8613% | 0.4891% | 0.2216% | 0.7272% | 1.28% | 0.8104% | 0.1911% | 0.7391% | 0.2301% | 2.293% | 0.6087% | 0.7643%
| PDM3 |  | 1.704% | -1.186% | -1.431% | 1.851% | -2.264% | -0.1572% | 0.9206% | 0.3176% | 0.7737% | 1.362% | 3.076% | -0.1082% | 3.517% | -2.475% | -2.803% | 0.3817% | 4.497% | -2.705% | 2.684% | -3.831% | -3.489% | -0.2551% | -2.705%
| Z-Score (Absolute) |  | 0.62 | 0.37 | 0.46 | 0.67 | 0.74 | 0.02 | 0.35 | 0.15 | 0.3 | 0.51 | 1.1 | 0 | 1.25 | 0.81 | 0.93 | 0.17 | 1.58 | 0.89 | 0.96 | 1.28 | 1.16 | 0.05 | 0.89
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Ag in OREAS 553b (Indicative Value 0.032 ppm) | [18 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 23 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 3
3
| 1 | 1 | < 100 | 0.03 | [15 empty] | 1
| 1 | 2 | < 100 | 0.04 | [15 empty] | 4
| 1 | 3 | < 100 | 0.027 | [15 empty] | 16
| 1 | 4 | < 100 | 0.026 | [15 empty] | 0.0315
| 1 | 5 | < 100 | 0.034 | [15 empty] | 10
| 1 | 6 | < 100 | 0.032
| Mean |  |  | 0.0315
| Median |  |  | 0.031
| Std Dev. |  |  | 0.0051284
| Rel.Std.Dev. |  |  | 16.28%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Al2O3 in OREAS 553b (Certified Value 12.51 wt.%) | [18 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 12.57 | 12.7 | 12.76 | 12.585 | 12.11 | 12.32 | 12.45 | 12.35 | 12.75 | 12.75 | 12.4485 | 12.32 | 12.7 | 13.53 | 11.27 | 12.4 |  | 1
| 1 | 2 | 12.56 | 12.43 | 12.73 | 12.663 | 12.12 | 12.301 | 12.45 | 12.55 | 12.55 | 12.7 | 12.3564 | 12.05 | 12.77 | 13.51 | 10.63 | 12.4 |  | #N/A
| 1 | 3 | 12.68 | 12.78 | 12.65 | 12.6 | 12.12 | 12.263 | 12.35 | 12.55 | 12.65 | 12.6 | 12.2243 | 11.6 | 12.77 | 13.55 | 12.05 | 12.4 |  | 16
| 1 | 4 | 12.51 | 12.43 | 12.73 | 12.56 | 11.96 | 12.338 | 12.55 | 12.45 | 12.55 | 12.75 | 12.304 | 12.38 | 12.79 | 13.46 | 10.7 | 12.4 |  | 12.5052
| 1 | 5 | 12.56 | 12.47 | 12.72 | 12.59 | 12.41 | 12.301 | 12.35 | 12.65 | 12.75 | 12.7 | 12.3703 | 12.58 | 12.76 | 13.41 | 11.17 | 12.3 |  | 73
| 1 | 6 | 12.56 | 12.66 | 12.76 | 12.6 | 12.42 | 12.168 | 12.45 | 12.35 | 12.7 | 12.7 | 12.3206 | 12.47 | 12.82 | 13.52 | 10.77 | 12.4
| Mean |  | 12.5733 | 12.5783 | 12.725 | 12.5997 | 12.19 | 12.2818 | 12.4333 | 12.4833 | 12.6583 | 12.7 | 12.3374 | 12.2333 | 12.7683 | 13.4967 | 11.0983 | 12.3833
| Median |  | 12.56 | 12.565 | 12.73 | 12.595 | 12.12 | 12.301 | 12.45 | 12.5 | 12.675 | 12.7 | 12.3385 | 12.35 | 12.77 | 13.515 | 10.97 | 12.4
| Std Dev. |  | 0.056451 | 0.15355 | 0.040373 | 0.034332 | 0.18461 | 0.061088 | 0.075277 | 0.12111 | 0.091742 | 0.054772 | 0.074806 | 0.35775 | 0.039707 | 0.052026 | 0.53406 | 0.040825
| Rel.Std.Dev. |  | 0.449% | 1.221% | 0.3173% | 0.2725% | 1.514% | 0.4974% | 0.6054% | 0.9701% | 0.7248% | 0.4313% | 0.6063% | 2.924% | 0.311% | 0.3855% | 4.812% | 0.3297%
| PDM3 |  | 0.5452% | 0.5852% | 1.758% | 0.7558% | -2.52% | -1.786% | -0.5743% | -0.1745% | 1.225% | 1.558% | -1.342% | -2.174% | 2.105% | 7.929% | -11.25% | -0.9741%
| Z-Score (Absolute) |  | 0.17 | 0.19 | 0.73 | 0.27 | 1.26 | 0.92 | 0.35 | 0.17 | 0.48 | 0.64 | 0.71 | 1.1 | 0.89 | 3.6 | 5.32 | 0.54
| Analytical results for As in OREAS 553b (Indicative Value 97 ppm) | [18 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 23 | 25 | 26 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 10 | < 100 | 200 | 100 | < 100 | 41 | < 80 | [10 empty] | 1
| 1 | 2 | < 10 | < 100 | 200 | 100 | < 100 | 44 | < 80 | [10 empty] | 5
| 1 | 3 | < 10 | < 100 | 200 | 100 | < 100 | 43 | < 80 | [10 empty] | 16
| 1 | 4 | < 10 | < 100 | 100 | 100 | < 100 | 43 | < 80 | [10 empty] | 97.1667
| 1 | 5 | < 10 | < 100 | 200 | < 100 | < 100 | 40 | < 80 | [10 empty] | 11
| 1 | 6 | < 10 | < 100 | 100 | < 100 | < 100 | 38 | < 80
| Mean |  |  |  | 166.6667 | 100 |  | 41.5
| Median |  |  |  | 200 | 100 |  | 42
| Std Dev. |  |  |  | 51.6398 | 0 |  | 2.2583
| Rel.Std.Dev. |  |  |  | 30.98% | 0% |  | 5.442%
| PDM3 |  |  |  | 71.53% | 2.916% |  | -57.29%
| Z-Score (Absolute) |  | 2.5 | 0 | 6.49 | 1.85 | 0 | 0.47 | 0.67
| Analytical results for BaO in OREAS 553b (Certified Value 303 ppm) | [18 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 22 | 25 | 26 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  |  |  |  | 0
0
| 1 | 1 | 300 | 260 | < 100 | < 11 | 300 | 300 | 400 | 400 | 332.2992 | 400 | 200 | 308.2 | 200 |  |  |  |  | 1
| 1 | 2 | 300 | 270 | < 100 | < 11 | 300 | 200 | 300 | 400 | 332.0222 | 300 | 200 | 289.2 | 300 |  |  |  |  | #N/A
| 1 | 3 | 300 | 300 | 100 | < 11 | 300 | 300 | 400 | 400 | 314.6381 | 300 | 200 | 297 | 300 |  |  |  |  | 16
| 1 | 4 | 300 | 300 | < 100 | < 11 | 300 | 200 | 300 | 400 | 297.1714 | 300 | 200 | 288.1 | 400 |  |  |  |  | 303.4666
| 1 | 5 | 300 | 300 | < 100 | < 11 | 300 | 300 | 400 | 300 | 301.5042 | 400 | 200 | 289.2 | 300 |  |  |  |  | 74
| 1 | 6 | 300 | 300 | 100 | < 11 | 300 | 300 | 300 | 400 | 327.7756 | 200 | 200 | 280.2 | 300
| Mean |  | 300 | 288.3333 | 100 |  | 300 | 266.6667 | 350 | 383.3333 | 317.5684 | 316.6667 | 200 | 291.9833 | 300
| Median |  | 300 | 300 | 100 |  | 300 | 300 | 350 | 400 | 321.2068 | 300 | 200 | 289.2 | 300
| Std Dev. |  | 0 | 18.3485 | 0 |  | 0 | 51.6398 | 54.7723 | 40.8248 | 15.5742 | 75.2773 | 0 | 9.5663 | 63.2456
| Rel.Std.Dev. |  | 0% | 6.364% | 0% |  | 0% | 19.36% | 15.65% | 10.65% | 4.904% | 23.77% | 0% | 3.276% | 21.08%
| PDM3 |  | -1.142% | -4.987% | -67.05% |  | -1.142% | -12.13% | 15.33% | 26.32% | 4.647% | 4.35% | -34.09% | -3.784% | -1.142%
| Z-Score (Absolute) |  | 0 | 0.45 | 8.96 | 11.3 | 0 | 1.28 | 1.92 | 3.2 | 0.67 | 0.64 | 3.84 | 0.31 | 0
| Analytical results for Be in OREAS 553b (Indicative Value 5 ppm) | [18 empty] | Indicative
Be | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 5 | [16 empty] | 1
| 1 | 2 | 6 | [16 empty] | 6
| 1 | 3 | 5 | [16 empty] | 16
| 1 | 4 | 5 | [16 empty] | 5
| 1 | 5 | 5 | [16 empty] | 12
| 1 | 6 | 4
| Mean |  | 5
| Median |  | 5
| Std Dev. |  | 0.63246
| Rel.Std.Dev. |  | 12.65%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 553b (Indicative Value < 100 ppm) | [18 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 23 | 25 | 26 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 10 | < 100 | 200 | < 100 | < 100 | 0.7 | < 90 | [10 empty] | 1
| 1 | 2 | < 10 | < 100 | 200 | < 100 | < 100 | 0.9 | < 90 | [10 empty] | 7
| 1 | 3 | < 10 | < 100 | 200 | < 100 | < 100 | 0.8 | < 90 | [10 empty] | 16
| 1 | 4 | < 10 | < 100 | 200 | < 100 | < 100 | 0.8 | < 90 | [10 empty] | < 100
| 1 | 5 | < 10 | < 100 | 200 | < 100 | < 100 | 0.7 | < 90 | [10 empty] | 13
| 1 | 6 | < 10 | < 100 | 200 | < 100 | < 100 | 0.7 | < 90
| Mean |  |  |  | 200 |  |  | 0.76667
| Median |  |  |  | 200 |  |  | 0.75
| Std Dev. |  |  |  | 0 |  |  | 0.08165
| Rel.Std.Dev. |  |  |  | 0% |  |  | 10.65%
PDM3
| Z-Score (Absolute) |  | 5.89 | 0 | 19.64 | 0 | 0 | 6.45 | 0.67
| Analytical results for CaO in OREAS 553b (Certified Value 0.435 wt.%) | [18 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.45 | 0.45 | 0.43 | 0.415 | 0.44 | 0.4477 | 0.42 | 0.41 | 0.45 | 0.44 | 0.411 | 0.38 | 0.44 | 0.56 | 0.35 | 0.45 |  | 1
| 1 | 2 | 0.44 | 0.46 | 0.43 | 0.415 | 0.44 | 0.4477 | 0.42 | 0.53 | 0.44 | 0.44 | 0.4028 | 0.36 | 0.44 | 0.55 | 0.34 | 0.45 |  | #N/A
| 1 | 3 | 0.44 | 0.46 | 0.42 | 0.43 | 0.45 | 0.4338 | 0.41 | 0.42 | 0.44 | 0.44 | 0.40291 | 0.35 | 0.45 | 0.57 | 0.38 | 0.45 |  | 16
| 1 | 4 | 0.44 | 0.46 | 0.43 | 0.42 | 0.43 | 0.4477 | 0.42 | 0.45 | 0.44 | 0.43 | 0.40905 | 0.4 | 0.44 | 0.5 | 0.32 | 0.44 |  | 0.43499
| 1 | 5 | 0.44 | 0.45 | 0.42 | 0.44 | 0.45 | 0.4477 | 0.41 | 0.42 | 0.45 | 0.43 | 0.409 | 0.38 | 0.45 | 0.54 | 0.34 | 0.44 |  | 75
| 1 | 6 | 0.43 | 0.45 | 0.43 | 0.43 | 0.46 | 0.4338 | 0.42 | 0.42 | 0.44 | 0.44 | 0.42217 | 0.38 | 0.45 | 0.5 | 0.31 | 0.44
| Mean |  | 0.44 | 0.455 | 0.42667 | 0.425 | 0.445 | 0.44307 | 0.41667 | 0.44167 | 0.44333 | 0.43667 | 0.40949 | 0.375 | 0.445 | 0.53667 | 0.34 | 0.445
| Median |  | 0.44 | 0.455 | 0.43 | 0.425 | 0.445 | 0.4477 | 0.42 | 0.42 | 0.44 | 0.44 | 0.40902 | 0.38 | 0.445 | 0.545 | 0.34 | 0.445
| Std Dev. |  | 0.0063246 | 0.0054772 | 0.005164 | 0.01 | 0.010488 | 0.0071779 | 0.005164 | 0.04535 | 0.005164 | 0.005164 | 0.0070954 | 0.017607 | 0.0054772 | 0.030111 | 0.024495 | 0.0054772
| Rel.Std.Dev. |  | 1.437% | 1.204% | 1.21% | 2.353% | 2.357% | 1.62% | 1.239% | 10.27% | 1.165% | 1.183% | 1.733% | 4.695% | 1.231% | 5.611% | 7.204% | 1.231%
| PDM3 |  | 1.151% | 4.6% | -1.914% | -2.297% | 2.301% | 1.856% | -4.213% | 1.534% | 1.917% | 0.3849% | -5.863% | -13.79% | 2.301% | 23.37% | -21.84% | 2.301%
| Z-Score (Absolute) |  | 0.06 | 1.04 | 1.04 | 1.16 | 0.31 | 0.17 | 1.78 | 0.06 | 0.18 | 0.31 | 2.31 | 4.84 | 0.31 | 7.05 | 7.42 | 0.31
| Analytical results for Cd in OREAS 553b (Indicative Value < 10 ppm) | [18 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 07 | 23 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | < 100 | [15 empty] | 1
| 1 | 2 | < 10 | < 100 | [15 empty] | 8
| 1 | 3 | < 10 | < 100 | [15 empty] | 16
| 1 | 4 | < 10 | < 100 | [15 empty] | < 10
| 1 | 5 | < 10 | < 100 | [15 empty] | 14
| 1 | 6 | < 10 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 553b (Indicative Value 151 ppm) | [18 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 160 | 160 | 149.61 | 162 | < 170 | [12 empty] | 1
| 1 | 2 | 240 | 160 | 149.59 | 146 | < 170 | [12 empty] | 9
| 1 | 3 | 160 | 160 | 149.62 | 160 | 260 | [12 empty] | 16
| 1 | 4 | 80 | 160 | 149.58 | 155 | 170 | [12 empty] | 151.3228
| 1 | 5 | 80 | 160 | 149.6 | 149 | < 170 | [12 empty] | 15
| 1 | 6 | 80 | 160 | 149.61 | 140 | 260
| Mean |  | 133.3333 | 160 | 149.6017 | 152 | 230
| Median |  | 120 | 160 | 149.605 | 152 | 260
| Std Dev. |  | 65.3197 | 0 | 0.01472 | 8.5088 | 51.9615
| Rel.Std.Dev. |  | 48.99% | 0% | 0.009839% | 5.598% | 22.59%
| PDM3 |  | -11.89% | 5.734% | -1.137% | 0.4475% | 51.99%
| Z-Score (Absolute) |  | 2.43 | 1.61 | 0.36 | 0 | 0.67
| Analytical results for Cl in OREAS 553b (Indicative Value 413 ppm) | [18 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 04 | 07 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 20 | 410 | [15 empty] | 1
| 1 | 2 | < 20 | 410 | [15 empty] | 10
| 1 | 3 | < 20 | 420 | [15 empty] | 16
| 1 | 4 | < 20 | 410 | [15 empty] | 413.3333
| 1 | 5 | < 20 | 410 | [15 empty] | 16
| 1 | 6 | < 20 | 420
| Mean |  |  | 413.3333
| Median |  |  | 410
| Std Dev. |  |  | 5.164
| Rel.Std.Dev. |  |  | 1.249%
| PDM3 |  |  | 8.882e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 553b (Certified Value 0.587 wt.%) | [18 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 13 | 14 | 19 | 20 | 22 | 23 | 25 | 26 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  |  | 3
3
| 1 | 1 | 0.59 | 0.5742 | 0.608 | 0.573 | 0.66 | 0.57 | 0.59 | 0.6 | 0.59 | 0.58 | 0.591 | 0.61 | 0.5912 | 0.55803 | 0.609 |  |  | 1
| 1 | 2 | 0.591 | 0.5565 | 0.606 | 0.58 | 0.66 | 0.57 | 0.59 | 0.59 | 0.6 | 0.56 | 0.593 | 0.6 | 0.5923 | 0.56863 | 0.609 |  |  | 30
| 1 | 3 | 0.591 | 0.5632 | 0.604 | 0.58 | 0.67 | 0.57 | 0.59 | 0.59 | 0.59 | 0.55 | 0.598 | 0.6 | 0.59 | 0.55046 | 0.617 |  |  | 16
| 1 | 4 | 0.584 | 0.5639 | 0.603 | 0.565 | 0.66 | 0.58 | 0.59 | 0.59 | 0.59 | 0.58 | 0.599 | 0.62 | 0.5837 | 0.55079 | 0.602 |  |  | 0.58666
| 1 | 5 | 0.585 | 0.569 | 0.612 | 0.587 | 0.66 | 0.57 | 0.6 | 0.6 | 0.59 | 0.59 | 0.597 | 0.61 | 0.5886 | 0.55286 | 0.609 |  |  | 76
| 1 | 6 | 0.588 | 0.5668 | 0.608 | 0.587 | 0.66 | 0.57 | 0.58 | 0.59 | 0.59 | 0.58 | 0.603 | 0.6 | 0.5734 | 0.55251 | 0.602
| Mean |  | 0.58817 | 0.5656 | 0.60683 | 0.57867 | 0.66167 | 0.57167 | 0.59 | 0.59333 | 0.59167 | 0.57333 | 0.59683 | 0.60667 | 0.58653 | 0.55555 | 0.608
| Median |  | 0.589 | 0.56535 | 0.607 | 0.58 | 0.66 | 0.57 | 0.59 | 0.59 | 0.59 | 0.58 | 0.5975 | 0.605 | 0.5893 | 0.55268 | 0.609
| Std Dev. |  | 0.0030605 | 0.0059736 | 0.0032506 | 0.008501 | 0.0040825 | 0.0040825 | 0.0063246 | 0.005164 | 0.0040825 | 0.015055 | 0.0043089 | 0.008165 | 0.0070969 | 0.006961 | 0.0055857
| Rel.Std.Dev. |  | 0.5203% | 1.056% | 0.5357% | 1.469% | 0.617% | 0.7141% | 1.072% | 0.8703% | 0.69% | 2.626% | 0.722% | 1.346% | 1.21% | 1.253% | 0.9187%
| PDM3 |  | 0.2576% | -3.589% | 3.439% | -1.362% | 12.79% | -2.555% | 0.5701% | 1.138% | 0.8542% | -2.271% | 1.735% | 3.411% | -0.02083% | -5.303% | 3.638%
| Z-Score (Absolute) |  | 0.07 | 0.99 | 0.68 | 0.45 | 2.9 | 0.74 | 0 | 0.13 | 0.07 | 0.67 | 0.28 | 0.67 | 0.14 | 1.39 | 0.74
| Analytical results for Cr2O3 in OREAS 553b (Indicative Value 143 ppm) | [18 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  |  | 0
0
| 1 | 1 | 200 | 200 | 120 | 400 | < 15 | 200 | 200 | 290 | 150 | 166.1496 | 60 | 120 | < 100 | 128.6 | 200 |  |  | 1
| 1 | 2 | 100 | 100 | 110 | 100 | < 15 | 100 | 200 | 290 | 150 | 166.0111 | 70 | 120 | < 100 | 118.4 | 200 |  |  | 11
| 1 | 3 | 200 | 100 | 120 | 100 | < 15 | 300 | 200 | 150 | 150 | 166.059 | 80 | 120 | < 100 | 134.5 | 200 |  |  | 16
| 1 | 4 | 100 | 100 | 70 | 100 | < 15 | 100 | 200 | 150 | 150 | 183.547 | 100 | 130 | < 100 | 116.9 | 200 |  |  | 143.353
| 1 | 5 | 200 | 200 | 110 | < 100 | < 15 | 100 | 200 | 150 | 150 | 161.6766 | 90 | 130 | < 100 | 112.5 | 100 |  |  | 17
| 1 | 6 | 200 | 100 | 100 | < 100 | < 15 | 100 | 200 | < 150 | 150 | 161.7026 | 130 | 130 | < 100 | 108.2 | 200
| Mean |  | 166.6667 | 133.3333 | 105 | 175 |  | 150 | 200 | 206 | 150 | 167.5243 | 88.3333 | 125 |  | 119.85 | 183.3333
| Median |  | 200 | 100 | 110 | 100 |  | 100 | 200 | 150 | 150 | 166.035 | 85 | 125 |  | 117.65 | 200
| Std Dev. |  | 51.6398 | 51.6398 | 18.7083 | 150 |  | 83.666 | 0 | 76.6812 | 0 | 8.1381 | 24.8328 | 5.4772 |  | 9.917 | 40.8248
| Rel.Std.Dev. |  | 30.98% | 38.73% | 17.82% | 85.71% |  | 55.78% | 0% | 37.22% | 0% | 4.858% | 28.11% | 4.382% |  | 8.275% | 22.27%
| PDM3 |  | 16.26% | -6.99% | -26.75% | 22.08% |  | 4.637% | 39.52% | 43.7% | 4.637% | 16.86% | -38.38% | -12.8% |  | -16.4% | 27.89%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.57 | 0 | 2.55 | 0.34 | 1.35 | 1 | 0.26 | 0.69 | 0.91 | 0.17 | 1.69 | 0.27 | 1.01
| Analytical results for Cs in OREAS 553b (Indicative Value 2.05 ppm) | [18 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 25 | 26 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [15 empty] | 2
2
| 1 | 1 | 2.1 | < 90 | [15 empty] | 1
| 1 | 2 | 2 | < 90 | [15 empty] | 12
| 1 | 3 | 2.1 | < 90 | [15 empty] | 16
| 1 | 4 | 2.1 | < 90 | [15 empty] | 2.05
| 1 | 5 | 2 | < 90 | [15 empty] | 18
| 1 | 6 | 2 | < 90
| Mean |  | 2.05
| Median |  | 2.05
| Std Dev. |  | 0.054772
| Rel.Std.Dev. |  | 2.672%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 553b (Certified Value 3.62 wt.%) | [18 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 3.632 | 3.6028 | 3.207 | 3.5491 | 3.731 | 3.56 | 3.62 | 3.65 | 3.66 | 3.57 | 3.6701 | 3.53 | 3.573 | 3.73 | 3.5072 | 2.0318 | 3.683 | 1
| 1 | 2 | 3.633 | 3.6268 | 3.576 | 3.5498 | 3.651 | 3.67 | 3.62 | 3.64 | 3.56 | 3.58 | 3.6767 | 3.42 | 3.55 | 3.74 | 3.5552 | 2.0238 | 3.723 | 31
| 1 | 3 | 3.609 | 3.6268 | 3.373 | 3.5389 | 3.675 | 3.69 | 3.6 | 3.67 | 3.56 | 3.58 | 3.6271 | 3.31 | 3.556 | 3.71 | 3.4913 | 2.2442 | 3.699 | 16
| 1 | 4 | 3.607 | 3.7067 | 3.129 | 3.5469 | 3.97 | 3.57 | 3.63 | 3.67 | 3.65 | 3.59 | 3.6307 | 3.54 | 3.581 | 3.73 | 3.6111 | 1.9411 | 3.707 | 3.6212
| 1 | 5 | 3.611 | 3.5868 | 3.49 | 3.5549 | 3.699 | 3.58 | 3.57 | 3.71 | 3.61 | 3.55 | 3.6665 | 3.61 | 3.608 | 3.71 | 3.5072 | 1.8412 | 3.683 | 77
| 1 | 6 | 3.607 | 3.6268 | 3.409 | 3.5469 | 3.826 | 3.6 | 3.59 | 3.64 | 3.72 | 3.53 | 3.6561 | 3.57 | 3.656 | 3.74 | 3.3155 | 1.9662 | 3.691
| Mean |  | 3.6165 | 3.6294 | 3.364 | 3.5477 | 3.7587 | 3.6117 | 3.605 | 3.6633 | 3.6267 | 3.5667 | 3.6545 | 3.4967 | 3.5873 | 3.7267 | 3.4979 | 2.008 | 3.6977
| Median |  | 3.61 | 3.6268 | 3.391 | 3.548 | 3.715 | 3.59 | 3.61 | 3.66 | 3.63 | 3.575 | 3.6613 | 3.535 | 3.577 | 3.73 | 3.5072 | 1.995 | 3.695
| Std Dev. |  | 0.012486 | 0.041274 | 0.169 | 0.005227 | 0.12009 | 0.054924 | 0.022583 | 0.026583 | 0.062503 | 0.022509 | 0.020984 | 0.1113 | 0.039404 | 0.013663 | 0.099582 | 0.13464 | 0.015526
| Rel.Std.Dev. |  | 0.3453% | 1.137% | 5.024% | 0.1473% | 3.195% | 1.521% | 0.6264% | 0.7257% | 1.723% | 0.6311% | 0.5742% | 3.183% | 1.098% | 0.3666% | 2.847% | 6.705% | 0.4199%
| PDM3 |  | -0.1289% | 0.2287% | -7.102% | -2.028% | 3.797% | -0.2624% | -0.4465% | 1.164% | 0.1518% | -1.505% | 0.9212% | -3.438% | -0.9344% | 2.913% | -3.404% | -44.55% | 2.113%
| Z-Score (Absolute) |  | 0.06 | 0.23 | 3.23 | 0.83 | 1.92 | 0 | 0.09 | 0.67 | 0.2 | 0.59 | 0.56 | 1.5 | 0.32 | 1.5 | 1.48 | 20.93 | 1.12
| Analytical results for Dy in OREAS 553b (Indicative Value 6.81 ppm) | [18 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 7.59 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 6.68 | < 170 | [14 empty] | 13
| 1 | 3 | < 10 | 7.51 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 6.41 | < 170 | [14 empty] | 6.81
| 1 | 5 | < 10 | 6.48 | < 170 | [14 empty] | 19
| 1 | 6 | < 10 | 6.19 | < 170
| Mean |  |  | 6.81
| Median |  |  | 6.58
| Std Dev. |  |  | 0.59474
| Rel.Std.Dev. |  |  | 8.733%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0 | 29.92
| Analytical results for Er in OREAS 553b (Indicative Value 4.08 ppm) | [18 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 4.37 | [15 empty] | 1
| 1 | 2 | < 10 | 4.11 | [15 empty] | 14
| 1 | 3 | < 10 | 4.52 | [15 empty] | 16
| 1 | 4 | < 10 | 3.96 | [15 empty] | 4.0833
| 1 | 5 | < 10 | 3.76 | [15 empty] | 20
| 1 | 6 | < 10 | 3.78
| Mean |  |  | 4.0833
| Median |  |  | 4.035
| Std Dev. |  |  | 0.31156
| Rel.Std.Dev. |  |  | 7.63%
| PDM3 |  |  | 6.661e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 553b (Indicative Value 0.94 ppm) | [18 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 1.01 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 0.94 | < 170 | [14 empty] | 15
| 1 | 3 | < 10 | 0.97 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 0.95 | < 170 | [14 empty] | 0.93667
| 1 | 5 | < 10 | 0.9 | < 170 | [14 empty] | 21
| 1 | 6 | < 10 | 0.85 | < 170
| Mean |  |  | 0.93667
| Median |  |  | 0.945
| Std Dev. |  |  | 0.055737
| Rel.Std.Dev. |  |  | 5.951%
| PDM3 |  |  | -3.331e-14%
| Z-Score (Absolute) |  | 0 | 0.67 | 13.5
| Analytical results for F in OREAS 553b (Indicative Value < 5000 ppm) | [18 empty] | Indicative
F | Round | Replicate | Lab | [16 empty] | 1
| No | No | 26 | [16 empty] | ppm
|  |  | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 5000 | [16 empty] | 1
| 1 | 2 | < 5000 | [16 empty] | 16
| 1 | 3 | < 5000 | [16 empty] | 16
| 1 | 4 | < 5000 | [16 empty] | < 5000
| 1 | 5 | < 5000 | [16 empty] | 22
| 1 | 6 | < 5000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 553b (Certified Value 4.6 wt.%) | [18 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 4.63 | 4.71 | 4.847 | 4.678 | 4.44 | 4.647 | 4.532 | 4.64 | 4.718 | 4.618 | 4.6226 | 4.56 | 4.56 | 3.94 | 4.5097 | 4.57 | 4.65 | 1
| 1 | 2 | 4.61 | 4.66 | 4.832 | 4.668 | 4.47 | 4.647 | 4.561 | 4.64 | 4.604 | 4.675 | 4.6294 | 4.41 | 4.56 | 4 | 4.6397 | 4.66 | 4.63 | #N/A
| 1 | 3 | 4.65 | 4.67 | 4.847 | 4.66 | 4.48 | 4.661 | 4.546 | 4.66 | 4.618 | 4.647 | 4.5748 | 4.31 | 4.61 | 3.9 | 4.4097 | 4.69 | 4.63 | 16
| 1 | 4 | 4.6 | 4.71 | 4.847 | 4.67 | 4.4 | 4.661 | 4.561 | 4.67 | 4.575 | 4.647 | 4.5697 | 4.57 | 4.61 | 3.95 | 4.5097 | 3.92 | 4.6 | 4.5996
| 1 | 5 | 4.6 | 4.66 | 4.875 | 4.68 | 4.56 | 4.618 | 4.475 | 4.71 | 4.689 | 4.604 | 4.6077 | 4.64 | 4.58 | 3.93 | 4.4697 | 4.44 | 4.59 | 78
| 1 | 6 | 4.63 | 4.7 | 4.861 | 4.67 | 4.52 | 4.618 | 4.504 | 4.63 | 4.604 | 4.604 | 4.5867 | 4.61 | 4.65 | 3.89 | 4.4697 | 4.5 | 4.61
| Mean |  | 4.62 | 4.685 | 4.8515 | 4.671 | 4.4783 | 4.642 | 4.5298 | 4.6583 | 4.6347 | 4.6325 | 4.5985 | 4.5167 | 4.595 | 3.935 | 4.5014 | 4.4633 | 4.6183
| Median |  | 4.62 | 4.685 | 4.847 | 4.67 | 4.475 | 4.647 | 4.539 | 4.65 | 4.611 | 4.6325 | 4.5972 | 4.565 | 4.595 | 3.935 | 4.4897 | 4.535 | 4.62
| Std Dev. |  | 0.02 | 0.02429 | 0.014721 | 0.0072388 | 0.056716 | 0.019616 | 0.034301 | 0.029269 | 0.055884 | 0.028501 | 0.025088 | 0.12863 | 0.035071 | 0.03937 | 0.077044 | 0.28232 | 0.022286
| Rel.Std.Dev. |  | 0.4329% | 0.5185% | 0.3034% | 0.155% | 1.266% | 0.4226% | 0.7572% | 0.6283% | 1.206% | 0.6152% | 0.5456% | 2.848% | 0.7633% | 1.001% | 1.712% | 6.325% | 0.4826%
| PDM3 |  | 0.4432% | 1.856% | 5.476% | 1.552% | -2.637% | 0.9215% | -1.517% | 1.277% | 0.7621% | 0.715% | -0.02483% | -1.803% | -0.1003% | -14.45% | -2.136% | -2.963% | 0.407%
| Z-Score (Absolute) |  | 0.02 | 0.85 | 2.98 | 0.67 | 1.79 | 0.3 | 1.13 | 0.51 | 0.21 | 0.18 | 0.25 | 1.3 | 0.3 | 8.75 | 1.5 | 1.98 | 0
| Analytical results for Ga in OREAS 553b (Indicative Value 27.3 ppm) | [18 empty] | Indicative
Ga | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 1
1
| 1 | 1 | 27.9 | [16 empty] | 1
| 1 | 2 | 26.7 | [16 empty] | 17
| 1 | 3 | 28.3 | [16 empty] | 16
| 1 | 4 | 27.3 | [16 empty] | 27.2667
| 1 | 5 | 27.2 | [16 empty] | 23
| 1 | 6 | 26.2
| Mean |  | 27.2667
| Median |  | 27.25
| Std Dev. |  | 0.76594
| Rel.Std.Dev. |  | 2.809%
| PDM3 |  | -1.443e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 553b (Indicative Value 7.01 ppm) | [18 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 7.44 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 6.94 | < 170 | [14 empty] | 18
| 1 | 3 | < 10 | 7.42 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 6.89 | < 170 | [14 empty] | 7.01
| 1 | 5 | < 10 | 6.44 | < 170 | [14 empty] | 24
| 1 | 6 | < 10 | 6.93 | < 170
| Mean |  |  | 7.01
| Median |  |  | 6.935
| Std Dev. |  |  | 0.37513
| Rel.Std.Dev. |  |  | 5.351%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0 | 26.75
| Analytical results for Hf in OREAS 553b (Indicative Value 42.5 ppm) | [18 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | < 80 | < 80 | < 10 | 7.5 | 80 | [12 empty] | 1
| 1 | 2 | < 80 | < 80 | < 10 | 7.1 | < 80 | [12 empty] | 22
| 1 | 3 | < 80 | < 80 | < 10 | 7.8 | 80 | [12 empty] | 16
| 1 | 4 | < 80 | < 80 | < 10 | 7.4 | 80 | [12 empty] | 42.4732
| 1 | 5 | < 80 | < 80 | < 10 | 6.7 | 80 | [12 empty] | 10
| 1 | 6 | < 80 | < 80 | < 10 | 6.8 | 80
| Mean |  |  |  |  | 7.2167 | 80
| Median |  |  |  |  | 7.25 | 80
| Std Dev. |  |  |  |  | 0.42622 | 0
| Rel.Std.Dev. |  |  |  |  | 5.906% | 0%
| PDM3 |  |  |  |  | -83.01% | 88.35%
| Z-Score (Absolute) |  | 0 | 0 | 0.72 | 0.67 | 0.68
| Analytical results for Ho in OREAS 553b (Indicative Value 1.38 ppm) | [18 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 1.42 | [15 empty] | 1
| 1 | 2 | < 10 | 1.33 | [15 empty] | 5
| 1 | 3 | < 10 | 1.5 | [15 empty] | 16
| 1 | 4 | < 10 | 1.36 | [15 empty] | 1.375
| 1 | 5 | < 10 | 1.32 | [15 empty] | 11
| 1 | 6 | < 10 | 1.32
| Mean |  |  | 1.375
| Median |  |  | 1.345
| Std Dev. |  |  | 0.072042
| Rel.Std.Dev. |  |  | 5.239%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 553b (Indicative Value 2.2 ppm) | [18 empty] | Indicative
In | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 2 | [16 empty] | 1
| 1 | 2 | 2 | [16 empty] | 6
| 1 | 3 | 2.2 | [16 empty] | 16
| 1 | 4 | 2.5 | [16 empty] | 2.2
| 1 | 5 | 2.4 | [16 empty] | 12
| 1 | 6 | 2.1
| Mean |  | 2.2
| Median |  | 2.15
| Std Dev. |  | 0.20976
| Rel.Std.Dev. |  | 9.535%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 553b (Certified Value 3.04 wt.%) | [18 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 3.11 | 2.99 | 3.01 | 3.106 | 2.92 | 2.9031 | 3.17 | 3.09 | 3.17 | 3.15 | 3.0869 | 3.016 | 3.093 | 3.74 | 2.79 | 3.04 |  | 1
| 1 | 2 | 3.11 | 2.97 | 3 | 3.104 | 2.96 | 2.9151 | 3.16 | 3.06 | 3.11 | 3.17 | 3.0581 | 2.923 | 3.101 | 3.7 | 2.76 | 3.04 |  | #N/A
| 1 | 3 | 3.05 | 2.96 | 3 | 3.1 | 2.97 | 2.9513 | 3.13 | 3.09 | 3.14 | 3.16 | 3.0293 | 2.838 | 3.1 | 3.66 | 3.05 | 3.04 |  | 16
| 1 | 4 | 3.09 | 3.02 | 3 | 3.11 | 2.85 | 2.9151 | 3.15 | 3.07 | 3.1 | 3.17 | 3.0547 | 3.086 | 3.106 | 3.72 | 2.92 | 3.05 |  | 3.043
| 1 | 5 | 3.09 | 2.98 | 3 | 3.12 | 3.01 | 2.9031 | 3.16 | 3.11 | 3.16 | 3.14 | 3.0671 | 3.069 | 3.107 | 3.65 | 2.89 | 3.03 |  | 79
| 1 | 6 | 3.09 | 2.97 | 2.99 | 3.1 | 2.99 | 2.9031 | 3.16 | 3.04 | 3.12 | 3.14 | 3.0575 | 3.053 | 3.111 | 3.71 | 2.71 | 3.04
| Mean |  | 3.09 | 2.9817 | 3 | 3.1067 | 2.95 | 2.9151 | 3.155 | 3.0767 | 3.1333 | 3.155 | 3.0589 | 2.9975 | 3.103 | 3.6967 | 2.8533 | 3.04
| Median |  | 3.09 | 2.975 | 3 | 3.105 | 2.965 | 2.9091 | 3.16 | 3.08 | 3.13 | 3.155 | 3.0578 | 3.0345 | 3.1035 | 3.705 | 2.84 | 3.04
| Std Dev. |  | 0.021909 | 0.02137 | 0.0063246 | 0.0075542 | 0.057619 | 0.018668 | 0.013784 | 0.025033 | 0.028048 | 0.013784 | 0.018696 | 0.097322 | 0.0063561 | 0.035024 | 0.12469 | 0.0063246
| Rel.Std.Dev. |  | 0.709% | 0.7167% | 0.2108% | 0.2432% | 1.953% | 0.6404% | 0.4369% | 0.8137% | 0.8951% | 0.4369% | 0.6112% | 3.247% | 0.2048% | 0.9474% | 4.37% | 0.208%
| PDM3 |  | 1.545% | -2.015% | -1.413% | 2.092% | -3.056% | -4.202% | 3.681% | 1.107% | 2.969% | 3.681% | 0.5235% | -1.495% | 1.972% | 21.48% | -6.233% | -0.09834%
| Z-Score (Absolute) |  | 0.22 | 0.84 | 0.66 | 0.38 | 1.15 | 1.49 | 0.85 | 0.09 | 0.64 | 0.85 | 0.09 | 0.69 | 0.34 | 6.14 | 2.09 | 0.27
| Analytical results for La in OREAS 553b (Indicative Value 81 ppm) | [18 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 90 | 90 | 79.16 | 82.5 | < 170 | [12 empty] | 1
| 1 | 2 | 90 | 90 | 79.34 | 74.5 | < 170 | [12 empty] | 7
| 1 | 3 | 90 | 90 | 78.94 | 81.2 | < 170 | [12 empty] | 16
| 1 | 4 | 90 | 90 | 79.35 | 75.6 | < 170 | [12 empty] | 81.4472
| 1 | 5 | 90 | 90 | 79.21 | 74.1 | < 170 | [12 empty] | 13
| 1 | 6 | 170 | 90 | 78.84 | 68.8 | < 170
| Mean |  | 103.3333 | 90 | 79.14 | 76.1167
| Median |  | 90 | 90 | 79.185 | 75.05
| Std Dev. |  | 32.6599 | 0 | 0.20948 | 5.0412
| Rel.Std.Dev. |  | 31.61% | 0% | 0.2647% | 6.623%
| PDM3 |  | 26.87% | 10.5% | -2.833% | -6.545%
| Z-Score (Absolute) |  | 1.56 | 0 | 0.67 | 1.01 | 0
| Analytical results for Lu in OREAS 553b (Indicative Value 0.55 ppm) | [18 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 0.62 | [15 empty] | 1
| 1 | 2 | < 10 | 0.52 | [15 empty] | 8
| 1 | 3 | < 10 | 0.55 | [15 empty] | 16
| 1 | 4 | < 10 | 0.54 | [15 empty] | 0.545
| 1 | 5 | < 10 | 0.52 | [15 empty] | 14
| 1 | 6 | < 10 | 0.52
| Mean |  |  | 0.545
| Median |  |  | 0.53
| Std Dev. |  |  | 0.038859
| Rel.Std.Dev. |  |  | 7.13%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 553b (Certified Value 5.94 wt.%) | [18 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 6.03 | 5.83 | 6.02 | 5.878 | 5.95 | 5.82 | 5.95 | 5.83 | 6.16 | 6.08 | 6.0391 | 5.81 | 6.13 | 5.46 | 5.73 | 6 |  | 1
| 1 | 2 | 5.96 | 5.82 | 6.04 | 5.859 | 5.9 | 5.754 | 5.95 | 5.81 | 6.04 | 6.11 | 6.0009 | 5.66 | 6.12 | 5.47 | 5.61 | 5.9 |  | #N/A
| 1 | 3 | 5.97 | 5.89 | 6.01 | 5.87 | 5.92 | 5.837 | 5.89 | 5.85 | 6.05 | 6.08 | 5.9239 | 5.46 | 6.17 | 5.7 | 6.1 | 5.9 |  | 16
| 1 | 4 | 5.96 | 5.85 | 5.99 | 5.88 | 5.79 | 5.771 | 6.02 | 5.84 | 6.08 | 6.08 | 5.974 | 5.84 | 6.17 | 5.6 | 5.72 | 6 |  | 5.9391
| 1 | 5 | 5.95 | 5.91 | 6.01 | 5.86 | 6.05 | 5.737 | 5.89 | 5.96 | 6.15 | 6.05 | 5.9995 | 5.93 | 6.15 | 5.55 | 5.67 | 5.9 |  | 80
| 1 | 6 | 5.96 | 5.82 | 6.01 | 5.87 | 6.04 | 5.82 | 5.94 | 5.8 | 6.12 | 6.06 | 5.9821 | 5.89 | 6.18 | 5.34 | 5.25 | 6
| Mean |  | 5.9717 | 5.8533 | 6.0133 | 5.8695 | 5.9417 | 5.7898 | 5.94 | 5.8483 | 6.1 | 6.0767 | 5.9866 | 5.765 | 6.1533 | 5.52 | 5.68 | 5.95
| Median |  | 5.96 | 5.84 | 6.01 | 5.87 | 5.935 | 5.7955 | 5.945 | 5.835 | 6.1 | 6.08 | 5.9908 | 5.825 | 6.16 | 5.51 | 5.695 | 5.95
| Std Dev. |  | 0.029269 | 0.038297 | 0.01633 | 0.0087579 | 0.096626 | 0.04117 | 0.048166 | 0.057764 | 0.05099 | 0.020656 | 0.038031 | 0.17581 | 0.024221 | 0.12506 | 0.27218 | 0.054772
| Rel.Std.Dev. |  | 0.4901% | 0.6543% | 0.2716% | 0.1492% | 1.626% | 0.7111% | 0.8109% | 0.9877% | 0.8359% | 0.3399% | 0.6353% | 3.05% | 0.3936% | 2.266% | 4.792% | 0.9205%
| PDM3 |  | 0.5486% | -1.444% | 1.25% | -1.172% | 0.04343% | -2.513% | 0.01537% | -1.528% | 2.709% | 2.317% | 0.7999% | -2.931% | 3.607% | -7.056% | -4.362% | 0.1837%
| Z-Score (Absolute) |  | 0.23 | 0.66 | 0.54 | 0.53 | 0.01 | 1.13 | 0.01 | 0.69 | 1.19 | 1.02 | 0.34 | 1.32 | 1.59 | 3.15 | 1.95 | 0.07
| Analytical results for MnO in OREAS 553b (Certified Value 0.017 wt.%) | [18 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.02 | 0.022 | 0.02 | 0.02 | < 0.0013 | 0.013 | 0.02 | 0.013 | 0.013 | 0.014866 | 0.014 | 0.016 | 0.02 | 0.01953 | 0.017 | 0.019 |  | 1
| 1 | 2 | 0.02 | 0.025 | 0.02 | 0.02 | < 0.0013 | 0.013 | 0.02 | 0.013 | 0.013 | 0.014854 | 0.014 | 0.016 | 0.02 | 0.01581 | 0.016 | 0.019 |  | #N/A
| 1 | 3 | 0.02 | 0.027 | 0.02 | 0.01 | < 0.0013 | 0.013 | 0.02 | 0.013 | 0.013 | 0.013984 | 0.01 | 0.016 | 0.02 | 0.01488 | 0.019 | 0.009 |  | 16
| 1 | 4 | 0.02 | 0.028 | 0.02 | 0.01 | < 0.0013 | 0.013 | 0.02 | 0.013 | 0.013 | 0.013985 | 0.015 | 0.016 | 0.02 | 0.01488 | 0.016 | 0.009 |  | 0.016648
| 1 | 5 | 0.02 | 0.022 | 0.02 | 0.02 | < 0.0013 | 0.013 | 0.02 | 0.013 | 0.013 | 0.01442 | 0.015 | 0.016 | 0.02 | 0.0186 | 0.016 | 0.009 |  | 81
| 1 | 6 | 0.02 | 0.023 | 0.02 | 0.01 | < 0.0013 | 0.013 | 0.02 | 0.013 | 0.013 | 0.013548 | 0.014 | 0.018 | 0.02 | 0.01581 | 0.016 | 0.019
| Mean |  | 0.02 | 0.0245 | 0.02 | 0.015 |  | 0.013 | 0.02 | 0.013 | 0.013 | 0.014276 | 0.013667 | 0.016333 | 0.02 | 0.016585 | 0.016667 | 0.014
| Median |  | 0.02 | 0.024 | 0.02 | 0.015 |  | 0.013 | 0.02 | 0.013 | 0.013 | 0.014202 | 0.014 | 0.016 | 0.02 | 0.01581 | 0.016 | 0.014
| Std Dev. |  | 0 | 0.0025884 | 0 | 0.0054772 |  | 0 | 0 | 0 | 0 | 0.00052964 | 0.0018619 | 0.0008165 | 0 | 0.0019874 | 0.0012111 | 0.0054772
| Rel.Std.Dev. |  | 0% | 10.57% | 0% | 36.51% |  | 0% | 0% | 0% | 0% | 3.71% | 13.62% | 4.999% | 0% | 11.98% | 7.266% | 39.12%
| PDM3 |  | 20.14% | 47.17% | 20.14% | -9.898% |  | -21.91% | 20.14% | -21.91% | -21.91% | -14.25% | -17.91% | -1.888% | 20.14% | -0.3756% | 0.1138% | -15.9%
| Z-Score (Absolute) |  | 1.06 | 2.16 | 1.06 | 0.16 | 3.68 | 0.67 | 1.06 | 0.67 | 0.67 | 0.34 | 0.49 | 0.16 | 1.06 | 0.22 | 0.24 | 0.42
| Analytical results for Mo in OREAS 553b (Indicative Value 24.4 ppm) | [18 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 25 | 26 | [11 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | 36 | < 50 | < 50 | < 50 | 15.7 | 1000 | [11 empty] | 1
| 1 | 2 | 32 | < 50 | < 50 | < 50 | 15.4 | 1130 | [11 empty] | 9
| 1 | 3 | 32 | < 50 | < 50 | < 50 | 15.7 | 1000 | [11 empty] | 16
| 1 | 4 | 33 | < 50 | < 50 | < 50 | 14.8 | 1070 | [11 empty] | 24.375
| 1 | 5 | 34 | < 50 | < 50 | < 50 | 15.3 | 1000 | [11 empty] | 15
| 1 | 6 | 34 | < 50 | < 50 | < 50 | 14.6 | 1070
| Mean |  | 33.5 |  |  |  | 15.25 | 1045
| Median |  | 33.5 |  |  |  | 15.35 | 1035
| Std Dev. |  | 1.5166 |  |  |  | 0.45935 | 53.9444
| Rel.Std.Dev. |  | 4.527% |  |  |  | 3.012% | 5.162%
| PDM3 |  | 37.44% |  |  |  | -37.44% | 4187%
| Z-Score (Absolute) |  | 1.35 | 0 | 0 | 0 | 1.55 | 161.71
| Analytical results for Na2O in OREAS 553b (Indicative Value 0.074 wt.%) | [18 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 15 | 19 | 22 | 25 | 26 | [8 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 3
3
| 1 | 1 | 0.07 | 0.09 | 0.09 | < 0.01 | < 0.01 | 0.04 | 0.06 | 0.05 | 0.1 | [8 empty] | 1
| 1 | 2 | 0.08 | 0.11 | 0.1 | < 0.01 | < 0.01 | 0.03 | 0.06 | 0.06 | < 0.1 | [8 empty] | 10
| 1 | 3 | 0.08 | 0.11 | 0.1 | < 0.01 | < 0.01 | 0.02 | 0.06 | 0.05 | < 0.1 | [8 empty] | 16
| 1 | 4 | 0.07 | 0.08 | 0.11 | < 0.01 | < 0.01 | 0.55 | 0.06 | 0.06 | < 0.1 | [8 empty] | 0.074048
| 1 | 5 | 0.08 | 0.08 | 0.11 | < 0.01 | < 0.01 | 0.03 | 0.06 | 0.06 | < 0.1 | [8 empty] | 16
| 1 | 6 | 0.07 | 0.09 | 0.12 | < 0.01 | < 0.01 | 0.03 | 0.06 | 0.05 | < 0.1
| Mean |  | 0.075 | 0.093333 | 0.105 |  |  | 0.11667 | 0.06 | 0.055 | 0.1
| Median |  | 0.075 | 0.09 | 0.105 |  |  | 0.03 | 0.06 | 0.055 | 0.1
| Std Dev. |  | 0.0054772 | 0.013663 | 0.010488 |  |  | 0.21238 | 0 | 0.0054772
| Rel.Std.Dev. |  | 7.303% | 14.64% | 9.989% |  |  | 182% | 0% | 9.959%
| PDM3 |  | 1.286% | 26.05% | 41.8% |  |  | 57.56% | -18.97% | -25.72% | 35.05%
| Z-Score (Absolute) |  | 0.3 | 0.67 | 0.91 | 1.11 | 1.11 | 1.15 | 0 | 0.1 | 0.03
| Analytical results for Nb in OREAS 553b (Indicative Value 44.6 ppm) | [18 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 10 | 13 | 14 | 22 | 25 | 26 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | < 50 | 70 | 50 | 26.16 | 26.1 | < 70 | [11 empty] | 1
| 1 | 2 | < 50 | 50 | 60 | 26.34 | 25.1 | < 70 | [11 empty] | 26
| 1 | 3 | < 50 | 80 | 50 | 26.28 | 26.7 | < 70 | [11 empty] | 16
| 1 | 4 | < 50 | 70 | 60 | 26.22 | 25.5 | < 70 | [11 empty] | 44.5787
| 1 | 5 | < 50 | 70 | 70 | 26.31 | 24.7 | < 70 | [11 empty] | 17
| 1 | 6 | < 50 | 70 | 60 | 26.18 | 24.3 | < 70
| Mean |  |  | 68.3333 | 58.3333 | 26.2483 | 25.4
| Median |  |  | 70 | 60 | 26.25 | 25.3
| Std Dev. |  |  | 9.8319 | 7.5277 | 0.072778 | 0.89219
| Rel.Std.Dev. |  |  | 14.39% | 12.9% | 0.2773% | 3.513%
| PDM3 |  |  | 53.29% | 30.85% | -41.12% | -43.02%
| Z-Score (Absolute) |  | 0.7 | 4.71 | 3.46 | 0.54 | 0.65 | 0.54
| Analytical results for Nd in OREAS 553b (Indicative Value 63 ppm) | [18 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | 64.66 | 67.4 | < 170 | [14 empty] | 1
| 1 | 2 | 64.14 | 59.9 | < 170 | [14 empty] | 12
| 1 | 3 | 64.58 | 66.3 | < 170 | [14 empty] | 16
| 1 | 4 | 65.05 | 62.5 | < 170 | [14 empty] | 63.2992
| 1 | 5 | 63.99 | 58.6 | < 170 | [14 empty] | 18
| 1 | 6 | 64.37 | 58.1 | < 170
| Mean |  | 64.465 | 62.1333
| Median |  | 64.475 | 61.2
| Std Dev. |  | 0.38308 | 3.9743
| Rel.Std.Dev. |  | 0.5942% | 6.396%
| PDM3 |  | 1.842% | -1.842%
| Z-Score (Absolute) |  | 0 | 0.67 | 6.15
| Analytical results for Ni in OREAS 553b (Certified Value 125 ppm) | [18 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 04 | 10 | 13 | 14 | 20 | 25 | 26 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 120 | 120 | 130 | 220 | 110 | 130 | 83 | 160 | [9 empty] | 1
| 1 | 2 | 120 | 130 | 70 | 210 | 110 | 130 | 86 | 80 | [9 empty] | 29
| 1 | 3 | 120 | 160 | 90 | 80 | 140 | 140 | 90 | 80 | [9 empty] | 16
| 1 | 4 | 120 | 110 | 130 | < 50 | 120 | 140 | 88 | 160 | [9 empty] | 125.2575
| 1 | 5 | 110 | 130 | 130 | < 50 | 120 | 130 | 94 | 80 | [9 empty] | 82
| 1 | 6 | 130 | 150 | 130 | < 50 | 130 | 150 | 93 | 160
| Mean |  | 120 | 133.3333 | 113.3333 | 170 | 121.6667 | 136.6667 | 89 | 120
| Median |  | 120 | 130 | 130 | 210 | 120 | 135 | 89 | 120
| Std Dev. |  | 6.3246 | 18.619 | 26.5832 | 78.1025 | 11.6905 | 8.165 | 4.1952 | 43.8178
| Rel.Std.Dev. |  | 5.27% | 13.96% | 23.46% | 45.94% | 9.609% | 5.974% | 4.714% | 36.51%
| PDM3 |  | -4.197% | 6.447% | -9.52% | 35.72% | -2.867% | 9.109% | -28.95% | -4.197%
| Z-Score (Absolute) |  | 0.07 | 0.97 | 0.38 | 1.45 | 0.18 | 1.2 | 2.02 | 0.07
| Analytical results for P2O5 in OREAS 553b (Certified Value 0.083 wt.%) | [18 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.09 | 0.08 | 0.088 | 0.091 | 0.0802 | < 0.0023 | 0.07 | 0.08 | 0.08 | 0.09 | 0.078965 | 0.068 | 0.088 | 0.08 | 0.12 | 0.09 |  | 1
| 1 | 2 | 0.087 | 0.09 | 0.087 | 0.091 | 0.0802 | < 0.0023 | 0.08 | 0.09 | 0.08 | 0.09 | 0.079816 | 0.071 | 0.086 | 0.08 | 0.1 | 0.09 |  | #N/A
| 1 | 3 | 0.09 | 0.09 | 0.087 | 0.08 | 0.0871 | < 0.0023 | 0.07 | 0.07 | 0.08 | 0.08 | 0.078004 | 0.064 | 0.089 | 0.07 | 0.11 | 0.09 |  | 16
| 1 | 4 | 0.088 | 0.09 | 0.087 | 0.09 | 0.0779 | < 0.0023 | 0.08 | 0.08 | 0.08 | 0.09 | 0.078007 | 0.074 | 0.089 | 0.07 | 0.13 | 0.09 |  | 0.082926
| 1 | 5 | 0.088 | 0.08 | 0.086 | 0.08 | 0.0825 | < 0.0023 | 0.08 | 0.09 | 0.08 | 0.08 | 0.078916 | 0.076 | 0.089 | 0.08 | 0.12 | 0.09 |  | 83
| 1 | 6 | 0.087 | 0.09 | 0.089 | 0.08 | 0.0802 | < 0.0023 | 0.07 | 0.08 | 0.08 | 0.09 | 0.083942 | 0.071 | 0.089 | 0.1 | 0.11 | 0.09
| Mean |  | 0.088333 | 0.086667 | 0.087333 | 0.085333 | 0.08135 |  | 0.075 | 0.081667 | 0.08 | 0.086667 | 0.079608 | 0.070667 | 0.088333 | 0.08 | 0.115 | 0.09
| Median |  | 0.088 | 0.09 | 0.087 | 0.085 | 0.0802 |  | 0.075 | 0.08 | 0.08 | 0.09 | 0.07894 | 0.071 | 0.089 | 0.08 | 0.115 | 0.09
| Std Dev. |  | 0.0013663 | 0.005164 | 0.0010328 | 0.0058538 | 0.0031703 |  | 0.0054772 | 0.0075277 | 0 | 0.005164 | 0.0022294 | 0.004274 | 0.0012111 | 0.010954 | 0.010488 | 1.5202e-17
| Rel.Std.Dev. |  | 1.547% | 5.958% | 1.183% | 6.86% | 3.897% |  | 7.303% | 9.218% | 0% | 5.958% | 2.8% | 6.048% | 1.371% | 13.69% | 9.12% | 1.689e-14%
| PDM3 |  | 6.521% | 4.511% | 5.315% | 2.903% | -1.9% |  | -9.558% | -1.518% | -3.528% | 4.511% | -4% | -14.78% | 6.521% | -3.528% | 38.68% | 8.531%
| Z-Score (Absolute) |  | 0.84 | 0.55 | 0.67 | 0.32 | 0.38 | 14.38 | 1.48 | 0.32 | 0.61 | 0.55 | 0.68 | 2.24 | 0.84 | 0.61 | 5.5 | 1.13
| Analytical results for Pb in OREAS 553b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 23 | 25 | 26 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 10 | < 50 | 90 | < 50 | < 100 | 10.6 | < 90 | [10 empty] | 1
| 1 | 2 | < 10 | < 50 | 70 | < 50 | < 100 | 10 | < 90 | [10 empty] | 13
| 1 | 3 | < 10 | < 50 | 80 | < 50 | < 100 | 10.5 | < 90 | [10 empty] | 16
| 1 | 4 | < 10 | < 50 | < 50 | 50 | < 100 | 12.3 | < 90 | [10 empty] | < 50
| 1 | 5 | < 10 | < 50 | 60 | < 50 | < 100 | 10 | < 90 | [10 empty] | 19
| 1 | 6 | < 10 | < 50 | < 50 | < 50 | < 100 | 9.4 | < 90
| Mean |  |  |  | 75 | 50 |  | 10.4667
| Median |  |  |  | 75 | 50 |  | 10.25
| Std Dev. |  |  |  | 12.9099 |  |  | 0.99532
| Rel.Std.Dev. |  |  |  | 17.21% |  |  | 9.509%
PDM3
| Z-Score (Absolute) |  | 0.87 | 0.15 | 1.05 | 0 | 0.75 | 0.67 | 0.62
| Analytical results for Pr in OREAS 553b (Indicative Value 18 ppm) | [18 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 1
1
| 1 | 1 | 18.26 | 19.4 | < 170 | [14 empty] | 1
| 1 | 2 | 18.19 | 17.6 | < 170 | [14 empty] | 14
| 1 | 3 | 18.34 | 19.6 | < 170 | [14 empty] | 16
| 1 | 4 | 18.22 | 17.5 | < 170 | [14 empty] | 18.0183
| 1 | 5 | 18.32 | 16.7 | < 170 | [14 empty] | 20
| 1 | 6 | 18.29 | 15.8 | < 170
| Mean |  | 18.27 | 17.7667
| Median |  | 18.275 | 17.55
| Std Dev. |  | 0.057966 | 1.4922
| Rel.Std.Dev. |  | 0.3173% | 8.399%
| PDM3 |  | 1.397% | -1.397%
| Z-Score (Absolute) |  | 0 | 0.67 | 86.41
| Analytical results for Rb in OREAS 553b (Indicative Value 67 ppm) | [18 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 13 | 14 | 25 | 26 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 80 | < 50 | 101 | < 180 | [13 empty] | 1
| 1 | 2 | 80 | < 50 | 95.6 | < 180 | [13 empty] | 15
| 1 | 3 | 110 | < 50 | 97.3 | < 180 | [13 empty] | 16
| 1 | 4 | 70 | 60 | 92.3 | < 180 | [13 empty] | 66.8556
| 1 | 5 | 80 | < 50 | 96.9 | < 180 | [13 empty] | 21
| 1 | 6 | 60 | < 50 | 90.3 | < 180
| Mean |  | 80 | 60 | 95.5667
| Median |  | 80 | 60 | 96.25
| Std Dev. |  | 16.7332 |  | 3.8135
| Rel.Std.Dev. |  | 20.92% |  | 3.99%
| PDM3 |  | 19.66% | -10.25% | 42.94%
| Z-Score (Absolute) |  | 0.5 | 4.76 | 0.85 | 0.5
| Analytical results for Re in OREAS 553b (Indicative Value < 0.1 ppm) | [18 empty] | Indicative
Re | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | < 0.1 | [16 empty] | 1
| 1 | 2 | < 0.1 | [16 empty] | 16
| 1 | 3 | < 0.1 | [16 empty] | 16
| 1 | 4 | < 0.1 | [16 empty] | < 0.1
| 1 | 5 | < 0.1 | [16 empty] | 22
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 553b (Indicative Value 30.4 ppm) | [18 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 07 | 10 | 13 | 14 | 25 | 26 | [11 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | < 10 | < 50 | 50 | < 50 | 3 | 80 | [11 empty] | 1
| 1 | 2 | < 10 | < 50 | 70 | < 50 | 3 | < 80 | [11 empty] | 17
| 1 | 3 | < 10 | < 50 | < 50 | < 50 | 3 | 80 | [11 empty] | 16
| 1 | 4 | < 10 | 50 | < 50 | < 50 | 3 | < 80 | [11 empty] | 30.3687
| 1 | 5 | < 10 | 60 | < 50 | < 50 | 3 | < 80 | [11 empty] | 23
| 1 | 6 | < 10 | < 50 | < 50 | < 50 | 3 | 170
| Mean |  |  | 55 | 60 |  | 3 | 110
| Median |  |  | 55 | 60 |  | 3 | 80
| Std Dev. |  |  | 7.0711 | 14.1421 |  | 0 | 51.9615
| Rel.Std.Dev. |  |  | 12.86% | 23.57% |  | 0% | 47.24%
| PDM3 |  |  | 81.11% | 97.57% |  | -90.12% | 262.2%
| Z-Score (Absolute) |  | 1.06 | 0.21 | 0.28 | 0.21 | 1.15 | 1.98
| Analytical results for Sc in OREAS 553b (Indicative Value 14.7 ppm) | [18 empty] | Indicative
Sc | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 1
1
| 1 | 1 | 10 | [16 empty] | 1
| 1 | 2 | 18 | [16 empty] | 18
| 1 | 3 | 18 | [16 empty] | 16
| 1 | 4 | 16 | [16 empty] | 14.6667
| 1 | 5 | 13 | [16 empty] | 24
| 1 | 6 | 13
| Mean |  | 14.6667
| Median |  | 14.5
| Std Dev. |  | 3.2042
| Rel.Std.Dev. |  | 21.85%
| PDM3 |  | -2.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 553b (Indicative Value 3.83 ppm) | [18 empty] | Indicative
Se | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 4 | [16 empty] | 1
| 1 | 2 | 3 | [16 empty] | 21
| 1 | 3 | 4 | [16 empty] | 16
| 1 | 4 | 4 | [16 empty] | 3.8333
| 1 | 5 | 5 | [16 empty] | 10
| 1 | 6 | 3
| Mean |  | 3.8333
| Median |  | 4
| Std Dev. |  | 0.75277
| Rel.Std.Dev. |  | 19.64%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 553b (Certified Value 53.72 wt.%) | [18 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 53.87 | 53.79 | 54.4 | 53.561 | 52.51 | 52.902 | 53.5 | 54.2 | 54.7 | 54 | 53.5137 | 54.51 | 53.81 | 45.63 | 116.2612 | 56.29 | 53.8 | 1
| 1 | 2 | 53.77 | 53.08 | 54.34 | 53.574 | 52.65 | 52.731 | 53.5 | 54.1 | 53.6 | 54.2 | 53.138 | 53.31 | 53.84 | 45.49 | 116.2163 | 56.81 | 53.7 | #N/A
| 1 | 3 | 53.86 | 53.04 | 53.97 | 53.58 | 52.72 | 53.523 | 53.1 | 53.9 | 54 | 54.1 | 52.5526 | 51.25 | 54.15 | 45.57 | 115.0932 | 58.05 | 53.2 | 16
| 1 | 4 | 53.72 | 53.46 | 54.22 | 53.56 | 51.43 | 53.138 | 53.9 | 54.1 | 53.8 | 54 | 52.9809 | 54.8 | 54.13 | 45.26 | 115.9468 | 58.16 | 53.4 | 53.7231
| 1 | 5 | 53.65 | 53.67 | 54.42 | 53.58 | 53.72 | 52.795 | 53 | 54.3 | 54.5 | 53.9 | 53.1406 | 55.6 | 53.94 | 45.56 | 116.2388 | 58.5 | 53.3 | 84
| 1 | 6 | 53.64 | 54.14 | 54.56 | 53.57 | 53.58 | 53.202 | 53.5 | 53.5 | 54.3 | 53.8 | 52.9382 | 55.38 | 54.25 | 45.71 | 116.98 | 54.72 | 53.6
| Mean |  | 53.7517 | 53.53 | 54.3183 | 53.5708 | 52.7683 | 53.0485 | 53.4167 | 54.0167 | 54.15 | 54 | 53.044 | 54.1417 | 54.02 | 45.5367 | 116.1227 | 57.0883 | 53.5
| Median |  | 53.745 | 53.565 | 54.37 | 53.572 | 52.685 | 53.02 | 53.5 | 54.1 | 54.15 | 54 | 53.0595 | 54.655 | 54.035 | 45.565 | 116.2276 | 57.43 | 53.5
| Std Dev. |  | 0.099883 | 0.42586 | 0.20341 | 0.0088638 | 0.8293 | 0.29774 | 0.32506 | 0.28577 | 0.42308 | 0.14142 | 0.31479 | 1.6299 | 0.18155 | 0.15436 | 0.61109 | 1.4415 | 0.23664
| Rel.Std.Dev. |  | 0.1858% | 0.7956% | 0.3745% | 0.01655% | 1.572% | 0.5613% | 0.6085% | 0.529% | 0.7813% | 0.2619% | 0.5934% | 3.01% | 0.3361% | 0.339% | 0.5262% | 2.525% | 0.4423%
| PDM3 |  | 0.05326% | -0.3594% | 1.108% | -0.2833% | -1.777% | -1.256% | -0.5703% | 0.5465% | 0.7947% | 0.5155% | -1.264% | 0.7792% | 0.5527% | -15.24% | 116.2% | 6.264% | -0.4152%
| Z-Score (Absolute) |  | 0 | 0.38 | 0.98 | 0.31 | 1.7 | 1.22 | 0.58 | 0.46 | 0.69 | 0.43 | 1.22 | 0.67 | 0.46 | 14.2 | 107.84 | 5.77 | 0.44
| Analytical results for Sm in OREAS 553b (Indicative Value 8.75 ppm) | [18 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 22 | 25 | 26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | 9.1 | < 170 | [14 empty] | 1
| 1 | 2 | < 10 | 8.8 | < 170 | [14 empty] | 5
| 1 | 3 | < 10 | 9.5 | < 170 | [14 empty] | 16
| 1 | 4 | < 10 | 8.7 | < 170 | [14 empty] | 8.75
| 1 | 5 | < 10 | 8.1 | < 170 | [14 empty] | 11
| 1 | 6 | < 10 | 8.3 | < 170
| Mean |  |  | 8.75
| Median |  |  | 8.75
| Std Dev. |  |  | 0.51284
| Rel.Std.Dev. |  |  | 5.861%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0 | 13.93
| Analytical results for Sn in OREAS 553b (Indicative Value 33 ppm) | [18 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 10 | 13 | 14 | 23 | 25 | 26 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 1
1
| 1 | 1 | < 50 | 60 | 60 | < 100 | 2.6 | < 80 | [11 empty] | 1
| 1 | 2 | < 50 | 50 | 80 | < 100 | 2.5 | < 80 | [11 empty] | 22
| 1 | 3 | < 50 | < 50 | 80 | < 100 | 2.9 | < 80 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | 50 | < 100 | 2.7 | < 80 | [11 empty] | 32.95
| 1 | 5 | < 50 | < 50 | 50 | < 100 | 2.5 | < 80 | [11 empty] | 12
| 1 | 6 | 50 | 50 | 70 | < 100 | 2.6 | < 80
| Mean |  | 50 | 53.3333 | 65 |  | 2.6333
| Median |  | 50 | 50 | 65 |  | 2.6
| Std Dev. |  |  | 5.7735 | 13.784 |  | 0.15055
| Rel.Std.Dev. |  |  | 10.83% | 21.21% |  | 5.717%
| PDM3 |  | 51.75% | 61.86% | 97.27% |  | -92.01%
| Z-Score (Absolute) |  | 0.65 | 0.01 | 1.67 | 0.69 | 2.37 | 0.01
| Analytical results for SO3 in OREAS 553b (Certified Value 8.59 wt.%) | [18 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 04 | 07 | 10 | 11 | 13 | 14 | 19 | 22 | 26 | [7 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [7 empty] | 2
3
| 1 | 1 | 8.43 | > 5 | 8.2907 | 8.69 | 8.59 | 8.515 | 8.64 | 9.013 | 8.79 | 8.6 | [7 empty] | 1
| 1 | 2 | 8.37 | > 5 | 8.3156 | 8.64 | 8.56 | 8.266 | 8.69 | 8.773 | 8.65 | 8.7 | [7 empty] | #N/A
| 1 | 3 | 8.16 | > 5 | 8.4405 | 8.59 | 8.59 | 8.366 | 8.765 | 8.502 | 8.67 | 8.6 | [7 empty] | 16
| 1 | 4 | 8.42 | > 5 | 8.2407 | 8.715 | 8.6 | 8.54 | 8.715 | 9.036 | 8.91 | 8.6 | [7 empty] | 8.5942
| 1 | 5 | 8.29 | > 5 | 8.3406 | 8.615 | 8.67 | 8.515 | 8.64 | 9.152 | 8.81 | 8.5 | [7 empty] | 85
| 1 | 6 | 8.31 | > 5 | 8.4405 | 8.615 | 8.54 | 8.49 | 8.59 | 9.12 | 8.86 | 8.6
| Mean |  | 8.33 |  | 8.3448 | 8.6442 | 8.5917 | 8.4487 | 8.6733 | 8.9327 | 8.7817 | 8.6
| Median |  | 8.34 |  | 8.3281 | 8.6275 | 8.59 | 8.5025 | 8.665 | 9.0245 | 8.8 | 8.6
| Std Dev. |  | 0.1006 |  | 0.08118 | 0.04852 | 0.04446 | 0.10868 | 0.062583 | 0.24952 | 0.10323 | 0.063246
| Rel.Std.Dev. |  | 1.208% |  | 0.9728% | 0.5613% | 0.5175% | 1.286% | 0.7216% | 2.793% | 1.176% | 0.7354%
| PDM3 |  | -3.074% |  | -2.902% | 0.5816% | -0.02925% | -1.693% | 0.921% | 3.939% | 2.182% | 0.06771%
| Z-Score (Absolute) |  | 1.2 | NA | 1.14 | 0.2 | 0.04 | 0.67 | 0.33 | 1.48 | 0.81 | 0
| Analytical results for SrO in OREAS 553b (Certified Value 96 ppm) | [18 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 10 | 11 | 13 | 14 | 15 | 19 | 22 | 25 | 26 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 120 | < 100 | 120 | 120 | < 50 | 90 | 100 | 81.6 | < 100 | [8 empty] | 1
| 1 | 2 | < 120 | < 100 | 120 | 120 | < 50 | 80 | 100 | 78.1 | < 100 | [8 empty] | #N/A
| 1 | 3 | < 120 | 100 | 120 | 120 | < 50 | 70 | 100 | 84 | < 100 | [8 empty] | 16
| 1 | 4 | < 120 | 100 | 120 | 120 | < 50 | 90 | 100 | 81.6 | < 100 | [8 empty] | 95.6006
| 1 | 5 | < 120 | 100 | 120 | 120 | < 50 | 90 | 100 | 79.2 | < 100 | [8 empty] | 86
| 1 | 6 | < 120 | < 100 | 120 | 120 | < 50 | 70 | 100 | 78.1 | < 100
| Mean |  |  | 100 | 120 | 120 |  | 81.6667 | 100 | 80.4333
| Median |  |  | 100 | 120 | 120 |  | 85 | 100 | 80.4
| Std Dev. |  |  | 0 | 0 | 0 |  | 9.8319 | 0 | 2.3602
| Rel.Std.Dev. |  |  | 0% | 0% | 0% |  | 12.04% | 0% | 2.934%
| PDM3 |  |  | 4.602% | 25.52% | 25.52% |  | -14.58% | 4.602% | -15.87%
| Z-Score (Absolute) |  | 0.67 | 0.17 | 1.2 | 1.2 | 1.76 | 0.04 | 0.62 | 0 | 0.96
| Analytical results for Ta in OREAS 553b (Indicative Value 69 ppm) | [18 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 10 | 25 | 26 | [14 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | 30 | 1.8 | 160 | [14 empty] | 1
| 1 | 2 | 30 | 1.8 | 160 | [14 empty] | 7
| 1 | 3 | 20 | 1.9 | 160 | [14 empty] | 16
| 1 | 4 | 30 | 1.9 | 160 | [14 empty] | 69.1952
| 1 | 5 | 30 | 1.8 | 160 | [14 empty] | 13
| 1 | 6 | 30 | 1.7 | 250
| Mean |  | 28.3333 | 1.8167 | 175
| Median |  | 30 | 1.8 | 160
| Std Dev. |  | 4.0825 | 0.075277 | 36.7423
| Rel.Std.Dev. |  | 14.41% | 4.144% | 21%
| PDM3 |  | -59.05% | -97.37% | 152.9%
| Z-Score (Absolute) |  | 0 | 0.67 | 3.79
| Analytical results for Tb in OREAS 553b (Indicative Value 1.2 ppm) | [18 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 1.23 | [15 empty] | 1
| 1 | 2 | < 10 | 1.19 | [15 empty] | 8
| 1 | 3 | < 10 | 1.27 | [15 empty] | 16
| 1 | 4 | < 10 | 1.16 | [15 empty] | 1.2
| 1 | 5 | < 10 | 1.13 | [15 empty] | 14
| 1 | 6 | < 10 | 1.22
| Mean |  |  | 1.2
| Median |  |  | 1.205
| Std Dev. |  |  | 0.050596
| Rel.Std.Dev. |  |  | 4.216%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 553b (Indicative Value 0.075 ppm) | [18 empty] | Indicative
Te | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 3
3
| 1 | 1 | 0.1 | [16 empty] | 1
| 1 | 2 | 0.1 | [16 empty] | 23
| 1 | 3 | 0.1 | [16 empty] | 16
| 1 | 4 | < 0.1 | [16 empty] | 0.075
| 1 | 5 | < 0.1 | [16 empty] | 15
| 1 | 6 | < 0.1
| Mean |  | 0.1
| Median |  | 0.1
| Std Dev. |  | 1.6997e-17
| Rel.Std.Dev. |  | 1.7e-14%
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 553b (Indicative Value 15.7 ppm) | [18 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 10 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [13 empty] | 1
1
| 1 | 1 | 20 | 16.35 | 16.3 | < 90 | [13 empty] | 1
| 1 | 2 | 20 | 16.36 | 15.3 | < 90 | [13 empty] | 10
| 1 | 3 | < 20 | 16.42 | 16.2 | < 90 | [13 empty] | 16
| 1 | 4 | < 20 | 16.41 | 15.6 | 180 | [13 empty] | 15.68
| 1 | 5 | < 20 | 16.43 | 15.4 | < 90 | [13 empty] | 16
| 1 | 6 | 20 | 16.37 | 15.1 | 260
| Mean |  | 20 | 16.39 | 15.65 | 220
| Median |  | 20 | 16.39 | 15.5 | 220
| Std Dev. |  | 0 | 0.034059 | 0.49295 | 56.5685
| Rel.Std.Dev. |  | 0% | 0.2078% | 3.15% | 25.71%
| PDM3 |  | 27.55% | 4.528% | -0.1913% | 1303%
| Z-Score (Absolute) |  | 0.99 | 0.36 | 0.36 | 83.93
| Analytical results for TiO2 in OREAS 553b (Certified Value 0.791 wt.%) | [18 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 13 | 14 | 15 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.8 | 0.8 | 0.73 | 0.826 | 0.75 | 0.8507 | 0.79 | 0.79 | 0.8 | 0.79 | 0.79304 | 0.79 | 0.82 | 0.79 | 0.769 | 0.78 |  | 1
| 1 | 2 | 0.8 | 0.81 | 0.74 | 0.816 | 0.77 | 0.8007 | 0.78 | 0.79 | 0.79 | 0.8 | 0.77057 | 0.78 | 0.81 | 0.8 | 0.725 | 0.79 |  | #N/A
| 1 | 3 | 0.79 | 0.8 | 0.72 | 0.84 | 0.77 | 0.784 | 0.76 | 0.79 | 0.8 | 0.8 | 0.75989 | 0.74 | 0.82 | 0.77 | 0.814 | 0.79 |  | 16
| 1 | 4 | 0.8 | 0.8 | 0.72 | 0.81 | 0.75 | 0.8007 | 0.79 | 0.8 | 0.78 | 0.8 | 0.78355 | 0.79 | 0.81 | 0.78 | 0.696 | 0.78 |  | 0.79148
| 1 | 5 | 0.79 | 0.79 | 0.73 | 0.82 | 0.78 | 0.8007 | 0.8 | 0.78 | 0.79 | 0.8 | 0.788 | 0.8 | 0.81 | 0.79 | 0.747 | 0.78 |  | 87
| 1 | 6 | 0.8 | 0.81 | 0.74 | 0.84 | 0.78 | 0.834 | 0.81 | 0.77 | 0.78 | 0.8 | 0.78449 | 0.8 | 0.81 | 0.8 | 0.701 | 0.78
| Mean |  | 0.79667 | 0.80167 | 0.73 | 0.82533 | 0.76667 | 0.8118 | 0.78833 | 0.78667 | 0.79 | 0.79833 | 0.77992 | 0.78333 | 0.81333 | 0.78833 | 0.742 | 0.78333
| Median |  | 0.8 | 0.8 | 0.73 | 0.823 | 0.77 | 0.8007 | 0.79 | 0.79 | 0.79 | 0.8 | 0.78402 | 0.79 | 0.81 | 0.79 | 0.736 | 0.78
| Std Dev. |  | 0.005164 | 0.0075277 | 0.0089443 | 0.012501 | 0.013663 | 0.025094 | 0.017224 | 0.010328 | 0.0089443 | 0.0040825 | 0.012332 | 0.022509 | 0.005164 | 0.01169 | 0.044775 | 0.005164
| Rel.Std.Dev. |  | 0.6482% | 0.939% | 1.225% | 1.515% | 1.782% | 3.091% | 2.185% | 1.313% | 1.132% | 0.5114% | 1.581% | 2.874% | 0.6349% | 1.483% | 6.034% | 0.6592%
| PDM3 |  | 0.6549% | 1.287% | -7.768% | 4.277% | -3.135% | 2.567% | -0.398% | -0.6086% | -0.1874% | 0.8655% | -1.46% | -1.03% | 2.761% | -0.398% | -6.252% | -1.03%
| Z-Score (Absolute) |  | 0.61 | 0.98 | 4.27 | 2.71 | 1.59 | 1.72 | 0 | 0.12 | 0.12 | 0.73 | 0.62 | 0.37 | 1.83 | 0 | 3.39 | 0.37
| Analytical results for Tl in OREAS 553b (Indicative Value 0.57 ppm) | [18 empty] | Indicative
Tl | Round | Replicate | Lab | [16 empty] | 1
| No | No | 25 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 0.61 | [16 empty] | 1
| 1 | 2 | 0.57 | [16 empty] | 27
| 1 | 3 | 0.63 | [16 empty] | 16
| 1 | 4 | 0.57 | [16 empty] | 0.57333
| 1 | 5 | 0.56 | [16 empty] | 17
| 1 | 6 | 0.5
| Mean |  | 0.57333
| Median |  | 0.57
| Std Dev. |  | 0.045019
| Rel.Std.Dev. |  | 7.852%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 553b (Indicative Value 0.55 ppm) | [18 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 0.6 | [15 empty] | 1
| 1 | 2 | < 10 | 0.56 | [15 empty] | 12
| 1 | 3 | < 10 | 0.59 | [15 empty] | 16
| 1 | 4 | < 10 | 0.53 | [15 empty] | 0.55333
| 1 | 5 | < 10 | 0.53 | [15 empty] | 18
| 1 | 6 | < 10 | 0.51
| Mean |  |  | 0.55333
| Median |  |  | 0.545
| Std Dev. |  |  | 0.036148
| Rel.Std.Dev. |  |  | 6.533%
| PDM3 |  |  | 6.661e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 553b (Indicative Value 46.7 ppm) | [18 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 10 | 25 | 26 | [14 empty] | ppm
|  |  | OxBF*XRF | BF*XRF | OxBF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | 9 | < 80 | [14 empty] | 1
| 1 | 2 | < 10 | 8.21 | < 80 | [14 empty] | 13
| 1 | 3 | < 10 | 8.85 | < 80 | [14 empty] | 16
| 1 | 4 | < 10 | 8.69 | 170 | [14 empty] | 46.7366
| 1 | 5 | < 10 | 8.58 | < 80 | [14 empty] | 19
| 1 | 6 | < 10 | 8.69 | 170
| Mean |  |  | 8.67 | 170
| Median |  |  | 8.69 | 170
| Std Dev. |  |  | 0.269 | 0
| Rel.Std.Dev. |  |  | 3.103% | 0%
| PDM3 |  |  | -81.45% | 263.7%
| Z-Score (Absolute) |  | 0.67 | 0 | 13.99
| Analytical results for V2O5 in OREAS 553b (Certified Value 614 ppm) | [18 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 06 | 10 | 13 | 14 | 19 | 25 | 26 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 610 | 600 | 540 | 540 | 540 | 740 | 673 | 700 | [9 empty] | 1
| 1 | 2 | 680 | 600 | 540 | 540 | 540 | 740 | 632 | 600 | [9 empty] | #N/A
| 1 | 3 | 640 | 600 | 540 | 540 | 540 | 700 | 683.7 | 700 | [9 empty] | 16
| 1 | 4 | 680 | 600 | 540 | 540 | 540 | 780 | 662.3 | 700 | [9 empty] | 614.1301
| 1 | 5 | 570 | 600 | 540 | 540 | 540 | 800 | 642.7 | 700 | [9 empty] | 88
| 1 | 6 | 640 | 600 | 540 | 540 | 540 | 670 | 594.5 | 700
| Mean |  | 636.6667 | 600 | 540 | 540 | 540 | 738.3333 | 648.0333 | 683.3333
| Median |  | 640 | 600 | 540 | 540 | 540 | 740 | 652.5 | 700
| Std Dev. |  | 42.269 | 0 | 0 | 0 | 0 | 48.3391 | 32.4047 | 40.8248
| Rel.Std.Dev. |  | 6.639% | 0% | 0% | 0% | 0% | 6.547% | 5% | 5.974%
| PDM3 |  | 3.67% | -2.301% | -12.07% | -12.07% | -12.07% | 20.22% | 5.521% | 11.27%
| Z-Score (Absolute) |  | 0.17 | 0.17 | 0.76 | 0.76 | 0.76 | 1.1 | 0.27 | 0.59
| Analytical results for W in OREAS 553b (Indicative Value 4.85 ppm) | [18 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 10 | 13 | 14 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 2
2
| 1 | 1 | < 8 | 20 | < 10 | 3.3 | < 80 | [12 empty] | 1
| 1 | 2 | 8 | < 10 | < 10 | 2.9 | < 80 | [12 empty] | 14
| 1 | 3 | < 8 | 10 | < 10 | 2.9 | < 80 | [12 empty] | 16
| 1 | 4 | < 8 | < 10 | < 10 | 2.9 | < 80 | [12 empty] | 4.8487
| 1 | 5 | < 8 | 10 | < 10 | 2.8 | < 80 | [12 empty] | 20
| 1 | 6 | < 8 | < 10 | < 10 | 3.1 | < 80
| Mean |  | 8 | 13.3333 |  | 2.9833
| Median |  | 8 | 10 |  | 2.9
| Std Dev. |  |  | 5.7735 |  | 0.18348
| Rel.Std.Dev. |  |  | 43.3% |  | 6.15%
| PDM3 |  | 64.99% | 175% |  | -38.47%
| Z-Score (Absolute) |  | 0.13 | 1.39 | 0 | 0.67 | 11.59
| Analytical results for Y in OREAS 553b (Indicative Value 49 ppm) | [18 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 13 | 14 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | 39 | 63 | < 10 | 37.4 | < 160 | [12 empty] | 1
| 1 | 2 | 55 | 71 | < 10 | 36.3 | < 160 | [12 empty] | 15
| 1 | 3 | 55 | 47 | < 10 | 38.8 | < 160 | [12 empty] | 16
| 1 | 4 | 47 | 55 | < 10 | 37.6 | < 160 | [12 empty] | 48.9983
| 1 | 5 | 63 | 63 | < 10 | 36.2 | < 160 | [12 empty] | 21
| 1 | 6 | 39 | 63 | < 10 | 34.2 | < 160
| Mean |  | 49.6667 | 60.3333 |  | 36.75
| Median |  | 51 | 63 |  | 36.85
| Std Dev. |  | 9.6885 | 8.2624 |  | 1.5719
| Rel.Std.Dev. |  | 19.51% | 13.69% |  | 4.277%
| PDM3 |  | 1.364% | 23.13% |  | -25%
| Z-Score (Absolute) |  | 0 | 0.54 | 2.31 | 0.67 | 1.48
| Analytical results for Yb in OREAS 553b (Indicative Value 3.75 ppm) | [18 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 22 | 25 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | 4 | [15 empty] | 1
| 1 | 2 | < 10 | 3.8 | [15 empty] | 16
| 1 | 3 | < 10 | 4.1 | [15 empty] | 16
| 1 | 4 | < 10 | 3.7 | [15 empty] | 3.75
| 1 | 5 | < 10 | 3.6 | [15 empty] | 22
| 1 | 6 | < 10 | 3.3
| Mean |  |  | 3.75
| Median |  |  | 3.75
| Std Dev. |  |  | 0.2881
| Rel.Std.Dev. |  |  | 7.683%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 553b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 07 | 10 | 13 | 14 | 20 | 23 | 25 | 26 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | 60 | < 50 | < 50 | 130 | < 50 | < 10 | < 100 | 16 | < 80 | [8 empty] | 1
| 1 | 2 | 60 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 13 | < 80 | [8 empty] | 17
| 1 | 3 | 60 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 16 | < 80 | [8 empty] | 16
| 1 | 4 | 60 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 16 | < 80 | [8 empty] | < 50
| 1 | 5 | 60 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | 14 | < 80 | [8 empty] | 23
| 1 | 6 | 60 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | 13 | < 80
| Mean |  | 60 |  |  | 130 |  | 20 |  | 14.6667
| Median |  | 60 |  |  | 130 |  | 20 |  | 15
| Std Dev. |  | 0 | [6 empty] | 1.5055
| Rel.Std.Dev. |  | 0% | [6 empty] | 10.27%
PDM3
| Z-Score (Absolute) |  | 1.56 | 0 | 0 | 0.78 | 0 | 0.78 | 1.11 | 0.46 | 0.67
| Analytical results for Zr in OREAS 553b (Certified Value 240 ppm) | [18 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 10 | 13 | 14 | 25 | 26 | [12 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 220 | 300 | 300 | 247 | 150 | [12 empty] | 1
| 1 | 2 | 220 | 300 | 300 | 238 | 150 | [12 empty] | 21
| 1 | 3 | 220 | 300 | 200 | 250 | 150 | [12 empty] | 16
| 1 | 4 | 220 | 300 | 200 | 226 | 220 | [12 empty] | 240.4354
| 1 | 5 | 220 | 300 | 200 | 256 | 220 | [12 empty] | 89
| 1 | 6 | 220 | 300 | 300 | 253 | 220
| Mean |  | 220 | 300 | 250 | 245 | 185
| Median |  | 220 | 300 | 250 | 248.5 | 185
| Std Dev. |  | 0 | 0 | 54.7723 | 11.1714 | 38.3406
| Rel.Std.Dev. |  | 0% | 0% | 21.91% | 4.56% | 20.72%
| PDM3 |  | -8.499% | 24.77% | 3.978% | 1.898% | -23.06%
| Z-Score (Absolute) |  | 0.67 | 1.62 | 0.15 | 0 | 1.76
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 553b (Certified Value 12.72 wt.%) | [22 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 19 | 20 | 22 | 23 | 25 | 26 | wt.%
|  |  | LOI*TGA | BF*OES | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | 2
3
| 1 | 1 | 12.21 | 12.52 | 12.8 | 13.01 | 13.2297 | 13.1917 | 12.74 | 12.9 | 13.6423 | 12.5 | 12.41 | 12.22 | 12.56 | 12.5528 | 12.68 | 12.92 | 12.4 | 12.77 | 12.74 | 13.61 | 12.66 | 1
| 1 | 2 | 12.3 | 12.61 | 12.78 | 13.09 | 13.285 | 13.0618 | 12.75 | 12.7 | 13.6186 | 12.42 | 12.4 | 12.19 | 12.47 | 12.6257 | 12.78 | 12.88 | 12.6 | 12.87 | 12.7 | 13.72 | 12.59 | #N/A
| 1 | 3 | 12.3 | 12.5 | 12.86 | 13.05 | 13.2895 | 12.8271 | 12.78 | 12.8 | 13.6255 | 12.37 | 12.44 | 12.1 | 12.43 | 12.6005 | 12.71 | 12.83 | 12.4 | 12.88 | 12.75 | 13.52 | 12.59 | 16
| 1 | 4 | 12.23 | 12.72 | 12.88 | 12.94 | 13.2768 | 13.1784 | 12.73 | 12.7 | 13.6386 | 12.32 | 12.5 | 12.13 | 12.52 | 12.5967 | 12.74 | 12.83 | 12.5 | 12.92 | 12.79 | 13.82 | 12.55 | 12.7221
| 1 | 5 | 12.19 | 12.49 | 12.69 | 12.97 | 13.1962 | 13.1904 | 12.78 | 12.9 | 13.6139 | 12.4 | 12.36 | 12.11 | 12.4 | 12.6071 | 12.89 | 12.87 | 12.4 | 12.86 | 12.85 | 13.82 | 12.63 | 91
| 1 | 6 | 12.3 | 12.61 | 12.71 | 12.94 | 13.1989 | 13.0712 | 12.82 | 12.7 | 13.6161 | 12.37 | 12.67 | 12.09 | 12.34 | 12.5932 | 12.77 | 12.75 | 12.5 | 12.79 | 12.6 | 15.45 | 12.61
| Mean |  | 12.255 | 12.575 | 12.7867 | 13 | 13.246 | 13.0868 | 12.7667 | 12.7833 | 13.6258 | 12.3967 | 12.4633 | 12.14 | 12.4533 | 12.596 | 12.7617 | 12.8467 | 12.4667 | 12.8483 | 12.7383 | 13.99 | 12.605
| Median |  | 12.265 | 12.565 | 12.79 | 12.99 | 13.2533 | 13.1248 | 12.765 | 12.75 | 13.622 | 12.385 | 12.425 | 12.12 | 12.45 | 12.5986 | 12.755 | 12.85 | 12.45 | 12.865 | 12.745 | 13.77 | 12.6
| Std Dev. |  | 0.050892 | 0.088713 | 0.076855 | 0.061319 | 0.043179 | 0.14031 | 0.033267 | 0.098319 | 0.012041 | 0.060882 | 0.11147 | 0.052915 | 0.080416 | 0.024078 | 0.073052 | 0.058195 | 0.08165 | 0.057067 | 0.084715 | 0.72493 | 0.037815
| Rel.Std.Dev. |  | 0.4153% | 0.7055% | 0.6011% | 0.4717% | 0.326% | 1.072% | 0.2606% | 0.7691% | 0.08837% | 0.4911% | 0.8944% | 0.4359% | 0.6457% | 0.1912% | 0.5724% | 0.453% | 0.6549% | 0.4442% | 0.665% | 5.182% | 0.3%
| PDM3 |  | -3.671% | -1.156% | 0.5078% | 2.185% | 4.119% | 2.867% | 0.3506% | 0.4816% | 7.104% | -2.558% | -2.034% | -4.575% | -2.112% | -0.9908% | 0.3113% | 0.9794% | -2.008% | 0.9925% | 0.1279% | 9.966% | -0.9202%
| Z-Score (Absolute) |  | 1.43 | 0.53 | 0.07 | 0.67 | 1.37 | 0.92 | 0.01 | 0.06 | 2.44 | 1.03 | 0.84 | 1.76 | 0.87 | 0.47 | 0 | 0.24 | 0.83 | 0.25 | 0.07 | 3.48 | 0.44
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 553b (Indicative Value 1.56 ppm) | [23 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 17 | 22 | 25 | 26 | [18 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*MS | [18 empty] | 2
2
| 1 | 1 | < 1 | 2.34 | 0.8 | < 10 | [18 empty] | 1
| 1 | 2 | < 1 | 2.32 | 0.8 | < 10 | [18 empty] | 4
| 1 | 3 | < 1 | 2.36 | 0.7 | < 10 | [18 empty] | 16
| 1 | 4 | < 1 | 2.35 | 0.8 | < 10 | [18 empty] | 1.5608
| 1 | 5 | < 1 | 2.32 | 0.8 | < 10 | [18 empty] | 26
| 1 | 6 | < 1 | 2.34 | 0.8 | < 10
| Mean |  |  | 2.3383 | 0.78333
| Median |  |  | 2.34 | 0.8
| Std Dev. |  |  | 0.016021 | 0.040825
| Rel.Std.Dev. |  |  | 0.6851% | 5.212%
| PDM3 |  |  | 49.81% | -49.81%
| Z-Score (Absolute) |  | 0.78 | 0.57 | 0.57 | 2.52
| Analytical results for Al in OREAS 553b (Certified Value 6.47 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 6.67 | 6.39 | 6.62 | 6.62 | 6.3839 | 6.44 | 6.52 | 6.07 | 6.406 | 6.43 | 6.457 | 6.563 | 6.298 | 6.695 | 6.4501 | 6.49 | 6.59 | 6.1594 | 6.51 | 0.61 | 4.96 | 6.55 | 1
| 1 | 2 | 6.69 | 6.4 | 6.6 | 6.59 | 6.3876 | 6.2 | 6.49 | 6.33 | 6.244 | 6.43 | 6.51 | 6.51 | 6.245 | 6.642 | 6.4226 | 6.38 | 6.75 | 6.2419 | 6.43 | 0.64 | 5.59 | 6.55 | #N/A
| 1 | 3 | 6.58 | 6.92 | 6.58 | 6.45 | 6.3939 | 6.3 | 6.52 | 6.1 | 6.4725 | 6.563 | 6.404 | 6.536 | 6.245 | 6.695 | 6.2989 | 6.38 | 6.57 | 6.6568 | 6.47 | 0.63 | 5.43 | 6.58 | 16
| 1 | 4 | 6.55 | 6.8 | 6.49 | 6.46 | 6.3876 | 6.32 | 6.51 | 6.31 | 6.3815 | 6.457 | 6.483 | 6.536 | 6.298 | 6.563 | 6.3591 | 6.46 | 6.4 | 6.242 | 6.41 | 0.62 | 5.55 | 6.44 | 6.4666
| 1 | 5 | 6.67 | 6.56 | 6.57 | 6.54 | 6.3859 | 6.18 | 6.52 | 6.06 | 6.2821 | 6.483 | 6.377 | 6.774 | 6.298 | 6.616 | 6.4316 | 6.44 | 6.52 | 6.4933 | 6.57 | 0.61 | 5.37 | 6.52 | 93
| 1 | 6 | 6.62 | 6.65 | 6.57 | 6.34 | 6.3915 | 6.6 | 6.54 | 6.23 | 6.3816 | 6.536 | 6.483 | 6.801 | 6.298 | 6.668 | 6.5156 | 6.56 | 6.5 | 6.1893 | 6.62 | 0.63 | 5.99 | 6.28
| Mean |  | 6.63 | 6.62 | 6.5717 | 6.5 | 6.3884 | 6.34 | 6.5167 | 6.1833 | 6.3613 | 6.4832 | 6.4523 | 6.62 | 6.2803 | 6.6465 | 6.413 | 6.4517 | 6.555 | 6.3305 | 6.5017 | 0.62333 | 5.4817 | 6.4867
| Median |  | 6.645 | 6.605 | 6.575 | 6.5 | 6.3876 | 6.31 | 6.52 | 6.165 | 6.3816 | 6.47 | 6.47 | 6.5495 | 6.298 | 6.655 | 6.4271 | 6.45 | 6.545 | 6.242 | 6.49 | 0.625 | 5.49 | 6.535
| Std Dev. |  | 0.056214 | 0.21364 | 0.04446 | 0.10373 | 0.0036891 | 0.15799 | 0.01633 | 0.12226 | 0.083933 | 0.055683 | 0.051457 | 0.1311 | 0.027369 | 0.051165 | 0.075202 | 0.068823 | 0.1164 | 0.19893 | 0.081588 | 0.012111 | 0.33529 | 0.11201
| Rel.Std.Dev. |  | 0.8479% | 3.227% | 0.6765% | 1.596% | 0.05775% | 2.492% | 0.2506% | 1.977% | 1.319% | 0.8589% | 0.7975% | 1.98% | 0.4358% | 0.7698% | 1.173% | 1.067% | 1.776% | 3.142% | 1.255% | 1.943% | 6.116% | 1.727%
| PDM3 |  | 2.527% | 2.372% | 1.625% | 0.5165% | -1.209% | -1.958% | 0.7743% | -4.38% | -1.628% | 0.2562% | -0.2206% | 2.372% | -2.88% | 2.782% | -0.8294% | -0.2309% | 1.367% | -2.105% | 0.5423% | -90.36% | -15.23% | 0.3103%
| Z-Score (Absolute) |  | 1.04 | 0.98 | 0.67 | 0.21 | 0.51 | 0.82 | 0.31 | 1.82 | 0.68 | 0.1 | 0.1 | 0.98 | 1.2 | 1.15 | 0.35 | 0.1 | 0.56 | 0.88 | 0.22 | 37.47 | 6.32 | 0.12
| Analytical results for As in OREAS 553b (Certified Value 53 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 22 | 25 | 26 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS |  |  |  |  | 0
0
| 1 | 1 | 53 | < 100 | 55 | 50 | 56 | 50 | 51.7979 | 100 | 100 | 100 | 100 | < 100 | 44 | 52 | < 100 | 47.2 | 55 | 59 |  |  |  |  | 1
| 1 | 2 | 55 | < 100 | 57 | 50 | 56 | 40 | 48.663 | 100 | 100 | 100 | 100 | < 100 | 44 | 51 | < 100 | 47.2 | 54 | 51 |  |  |  |  | 5
| 1 | 3 | 55 | < 100 | 47 | 50 | 56 | 70 | 50.8296 | 100 | < 100 | 100 | 100 | < 100 | 44 | 53 | < 100 | 47.3 | 57 | 66 |  |  |  |  | 16
| 1 | 4 | 53 | < 100 | 59 | 50 | 56 | 50 | 47.6215 | 100 | < 100 | 100 | 200 | < 100 | 44 | 53 | < 100 | 47.3 | 56 | 62 |  |  |  |  | 52.866
| 1 | 5 | 54 | < 100 | 62 | 50 | 56 | 70 | 53.9359 | 100 | 100 | < 100 | 100 | < 100 | 46 | 55 | < 100 | 47.2 | 58 | 70 |  |  |  |  | 94
| 1 | 6 | 57 | < 100 | 49 | 50 | 56 | 40 | 54.9066 | 100 | 100 | 100 | 100 | < 100 | 44 | 54 | 100 | 47.2 | 58 | 56
| Mean |  | 54.5 |  | 54.8333 | 50 | 56 | 53.3333 | 51.2924 | 100 | 100 | 100 | 116.6667 |  | 44.3333 | 53 | 100 | 47.2333 | 56.3333 | 60.6667
| Median |  | 54.5 |  | 56 | 50 | 56 | 50 | 51.3138 | 100 | 100 | 100 | 100 |  | 44 | 53 | 100 | 47.2 | 56.5 | 60.5
| Std Dev. |  | 1.5166 |  | 5.8109 | 0 | 0 | 13.6626 | 2.8604 | 0 | 0 | 0 | 40.8248 |  | 0.8165 | 1.4142 |  | 0.05164 | 1.633 | 6.8605
| Rel.Std.Dev. |  | 2.783% |  | 10.6% | 0% | 0% | 25.62% | 5.577% | 0% | 0% | 0% | 34.99% |  | 1.842% | 2.668% |  | 0.1093% | 2.899% | 11.31%
| PDM3 |  | 3.091% |  | 3.721% | -5.421% | 5.928% | 0.884% | -2.977% | 89.16% | 89.16% | 89.16% | 120.7% |  | -16.14% | 0.2535% | 89.16% | -10.65% | 6.559% | 14.76%
| Z-Score (Absolute) |  | 0.02 | 0.67 | 0.02 | 0.67 | 0.19 | 0.19 | 0.49 | 6.55 | 4.14 | 5.35 | 8.96 | 0.67 | 1.49 | 0.24 | 0.53 | 1.07 | 0.24 | 0.87
| Analytical results for B in OREAS 553b (Certified Value 217 ppm) | [23 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 08 | 09 | 19 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | [15 empty] | 0
0
| 1 | 1 | 230 | 230 | 220 | 212.6581 | 234 | 37 | 198 | [15 empty] | 1
| 1 | 2 | 229 | 220 | 220 | 203.2779 | 232 | 31 | 191 | [15 empty] | 1
| 1 | 3 | 230 | 220 | 210 | 201.0978 | 243 | 31 | 200 | [15 empty] | 16
| 1 | 4 | 228 | 230 | 230 | 198.8685 | 227 | 32 | 183 | [15 empty] | 217.1039
| 1 | 5 | 227 | 220 | 210 | 217.0051 | 240 | 29 | 188 | [15 empty] | 95
| 1 | 6 | 229 | 210 | 220 | 215.8331 | 222 | 24 | 196
| Mean |  | 228.8333 | 221.6667 | 218.3333 | 208.1234 | 233 | 30.6667 | 192.6667
| Median |  | 229 | 220 | 220 | 207.968 | 233 | 31 | 193.5
| Std Dev. |  | 1.169 | 7.5277 | 7.5277 | 7.9672 | 7.8486 | 4.2269 | 6.5013
| Rel.Std.Dev. |  | 0.5109% | 3.396% | 3.448% | 3.828% | 3.368% | 13.78% | 3.374%
| PDM3 |  | 5.403% | 2.102% | 0.5663% | -4.136% | 7.322% | -85.87% | -11.26%
| Z-Score (Absolute) |  | 0.67 | 0.21 | 0 | 0.66 | 0.94 | 12.05 | 1.65
| Analytical results for Ba in OREAS 553b (Certified Value 264 ppm) | [23 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [7 empty] | 0
0
| 1 | 1 | 259 | 269 | 270 | 271 | 274 | 281 | 240 | 261.696 | 264.2742 | 265 | 272 | 251 | 49.4 | 237 | 265 | [7 empty] | 1
| 1 | 2 | 263 | 260 | 270 | 269 | 244 | 278 | 260 | 257.0159 | 273.3558 | 263 | 270 | 256 | 49.8 | 247 | 256 | [7 empty] | #N/A
| 1 | 3 | 259 | 251 | 278 | 271 | 260 | 279 | 250 | 263.5226 | 261.7402 | 263 | 267 | 270 | 50.5 | 235 | 272 | [7 empty] | 16
| 1 | 4 | 267 | 251 | 272 | 270 | 235 | 280 | 260 | 254.8993 | 265.8722 | 262 | 264 | 252 | 49.6 | 236 | 284 | [7 empty] | 264.0065
| 1 | 5 | 257 | 260 | 270 | 272 | 224 | 280 | 250 | 257.8097 | 274.5312 | 264 | 277 | 262 | 50.3 | 216 | 273 | [7 empty] | 96
| 1 | 6 | 275 | 260 | 263 | 275 | 258 | 276 | 240 | 259.0478 | 264.0547 | 267 | 265 | 250 | 49.6 | 186 | 265
| Mean |  | 263.3333 | 258.5 | 270.5 | 271.3333 | 249.1667 | 279 | 250 | 258.9986 | 267.3047 | 264 | 269.1667 | 256.8333 | 49.8667 | 226.1667 | 269.1667
| Median |  | 261 | 260 | 270 | 271 | 251 | 279.5 | 250 | 258.4288 | 265.0732 | 263.5 | 268.5 | 254 | 49.7 | 235.5 | 268.5
| Std Dev. |  | 6.7429 | 6.775 | 4.8062 | 2.0656 | 18.2912 | 1.7889 | 8.9443 | 3.1589 | 5.3217 | 1.7889 | 4.8751 | 7.8081 | 0.43665 | 22.1035 | 9.4956
| Rel.Std.Dev. |  | 2.561% | 2.621% | 1.777% | 0.7613% | 7.341% | 0.6412% | 3.578% | 1.22% | 1.991% | 0.6776% | 1.811% | 3.04% | 0.8756% | 9.773% | 3.528%
| PDM3 |  | -0.255% | -2.086% | 2.46% | 2.775% | -5.621% | 5.679% | -5.305% | -1.897% | 1.249% | -0.00245% | 1.955% | -2.717% | -81.11% | -14.33% | 1.955%
| Z-Score (Absolute) |  | 0 | 0.53 | 0.74 | 0.83 | 1.47 | 1.63 | 1.38 | 0.45 | 0.41 | 0.07 | 0.61 | 0.67 | 22.14 | 3.86 | 0.61
| Analytical results for Be in OREAS 553b (Certified Value 5.42 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 17 | 18 | 22 | 25 | 26 | [9 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | [9 empty] | 2
2
| 1 | 1 | 6 | 5.3 | 7 | 5 | 6 | 7 | 6 | 5.1705 | 6 | 6 | 3.1 | 6 | 5.5 | [9 empty] | 1
| 1 | 2 | 6 | 4.7 | 6 | 4 | 6 | 7 | 6 | 5.064 | 6 | 6 | 3.1 | 5 | 5.1 | [9 empty] | 6
| 1 | 3 | 6 | 2.4 | 6 | 4 | 5 | 7 | 5 | 5.0739 | 6 | 5 | 3.1 | 6 | 5.4 | [9 empty] | 16
| 1 | 4 | 5 | 5.7 | 6 | 4 | 5 | 7 | 6 | 4.9381 | 6 | 6 | 3.1 | 5 | 5.2 | [9 empty] | 5.4194
| 1 | 5 | 5 | 3.2 | 6 | 5 | 5 | 7 | 6 | 5.3652 | 6 | 5 | 3.1 | 6 | 5.3 | [9 empty] | 97
| 1 | 6 | 6 | 5.4 | 6 | 4 | 6 | 7 | 5 | 5.3113 | 6 | 6 | 3.1 | 4 | 5.1
| Mean |  | 5.6667 | 4.45 | 6.1667 | 4.3333 | 5.5 | 7 | 5.6667 | 5.1538 | 6 | 5.6667 | 3.1 | 5.3333 | 5.2667
| Median |  | 6 | 5 | 6 | 4 | 5.5 | 7 | 6 | 5.1222 | 6 | 6 | 3.1 | 5.5 | 5.25
| Std Dev. |  | 0.5164 | 1.3428 | 0.40825 | 0.5164 | 0.54772 | 0 | 0.5164 | 0.16174 | 0 | 0.5164 | 0 | 0.8165 | 0.1633
| Rel.Std.Dev. |  | 9.113% | 30.17% | 6.62% | 11.92% | 9.959% | 0% | 9.113% | 3.138% | 0% | 9.113% | 0% | 15.31% | 3.101%
| PDM3 |  | 4.562% | -17.89% | 13.79% | -20.04% | 1.486% | 29.16% | 4.562% | -4.901% | 10.71% | 4.562% | -42.8% | -1.589% | -2.819%
| Z-Score (Absolute) |  | 0.32 | 2.05 | 1.3 | 2.27 | 0 | 2.92 | 0.32 | 0.67 | 0.97 | 0.32 | 4.68 | 0.32 | 0.45
| Analytical results for Bi in OREAS 553b (Certified Value 1.77 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 17 | 18 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 1.9 | < 2 | 1.7 | 1.4405 | 1.9 | 1.8 | < 5 | 1.8 | 1.8 | [13 empty] | 1
| 1 | 2 | 1.8 | < 2 | 1.7 | 1.4674 | 1.8 | 1.7 | < 5 | 1.7 | 1.8 | [13 empty] | 7
| 1 | 3 | 2 | < 2 | 1.7 | 1.4167 | 1.8 | 1.6 | < 5 | 1.8 | 1.8 | [13 empty] | 16
| 1 | 4 | 1.8 | < 2 | 1.8 | 1.3855 | 1.8 | 1.6 | < 5 | 1.8 | 1.8 | [13 empty] | 1.7722
| 1 | 5 | 2 | < 2 | 1.6 | 1.3952 | 1.9 | 1.6 | < 5 | 1.8 | 1.8 | [13 empty] | 98
| 1 | 6 | 1.9 | < 2 | 1.6 | 1.4626 | 1.9 | 1.8 | < 5 | 1.5 | 1.7
| Mean |  | 1.9 |  | 1.6833 | 1.428 | 1.85 | 1.6833 |  | 1.7333 | 1.7833
| Median |  | 1.9 |  | 1.7 | 1.4286 | 1.85 | 1.65 |  | 1.8 | 1.8
| Std Dev. |  | 0.089443 |  | 0.075277 | 0.034364 | 0.054772 | 0.098319 |  | 0.12111 | 0.040825
| Rel.Std.Dev. |  | 4.708% |  | 4.472% | 2.406% | 2.961% | 5.841% |  | 6.987% | 2.289%
| PDM3 |  | 7.21% |  | -5.016% | -19.42% | 4.389% | -5.016% |  | -2.194% | 0.627%
| Z-Score (Absolute) |  | 0.96 | 4.24 | 0.29 | 1.76 | 0.67 | 0.29 | 4.43 | 0 | 0.29
| Analytical results for Ca in OREAS 553b (Certified Value 0.307 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.3 | 0.329 | 0.35 | 0.3 | 0.31223 | 0.5 | 0.25 | 0.3 | 0.34116 | 0.214 | 0.272 | 0.307 | 0.343 | 0.329 | 0.4 | 0.4 | 0.3503 | 0.3 | 0.3 | 0.24 | 0.4 |  | 1
| 1 | 2 | 0.3 | 0.329 | 0.32 | 0.3 | 0.31241 | 0.4 | 0.25 | 0.32 | 0.30501 | 0.214 | 0.257 | 0.322 | 0.372 | 0.329 | 0.3 | 0.4 | 0.3313 | 0.31 | 0.29 | 0.27 | 0.4 |  | #N/A
| 1 | 3 | 0.3 | 0.322 | 0.33 | 0.3 | 0.31217 | 0.4 | 0.24 | 0.3 | 0.34911 | 0.214 | 0.257 | 0.322 | 0.379 | 0.329 | 0.4 | 0.4 | 0.3686 | 0.32 | 0.28 | 0.23 | 0.4 |  | 16
| 1 | 4 | 0.3 | 0.336 | 0.33 | 0.3 | 0.3127 | 0.4 | 0.24 | 0.32 | 0.30949 | 0.229 | 0.257 | 0.329 | 0.357 | 0.329 | 0.3 | 0.4 | 0.3376 | 0.3 | 0.3 | 0.21 | 0.4 |  | 0.3067
| 1 | 5 | 0.3 | 0.336 | 0.34 | 0.3 | 0.3125 | 0.4 | 0.25 | 0.31 | 0.32805 | 0.272 | 0.257 | 0.293 | 0.329 | 0.329 | 0.3 | 0.4 | 0.348 | 0.32 | 0.29 | 0.2 | 0.4 |  | 99
| 1 | 6 | 0.3 | 0.336 | 0.21 | 0.3 | 0.31246 | 0.5 | 0.25 | 0.31 | 0.32999 | 0.257 | 0.272 | 0.293 | 0.3 | 0.343 | 0.3 | 0.4 | 0.3437 | 0.32 | 0.28 | 0.21 | 0.3
| Mean |  | 0.3 | 0.33133 | 0.31333 | 0.3 | 0.31241 | 0.43333 | 0.24667 | 0.31 | 0.32714 | 0.23333 | 0.262 | 0.311 | 0.34667 | 0.33133 | 0.33333 | 0.4 | 0.34658 | 0.31167 | 0.29 | 0.22667 | 0.38333
| Median |  | 0.3 | 0.3325 | 0.33 | 0.3 | 0.31243 | 0.4 | 0.25 | 0.31 | 0.32902 | 0.2215 | 0.257 | 0.3145 | 0.35 | 0.329 | 0.3 | 0.4 | 0.34585 | 0.315 | 0.29 | 0.22 | 0.4
| Std Dev. |  | 0 | 0.0057155 | 0.05164 | 0 | 0.00019113 | 0.05164 | 0.005164 | 0.0089443 | 0.017258 | 0.02528 | 0.007746 | 0.015684 | 0.029317 | 0.0057155 | 0.05164 | 6.0809e-17 | 0.012831 | 0.0098319 | 0.0089443 | 0.02582 | 0.040825
| Rel.Std.Dev. |  | 0% | 1.725% | 16.48% | 0% | 0.06118% | 11.92% | 2.094% | 2.885% | 5.275% | 10.83% | 2.956% | 5.043% | 8.457% | 1.725% | 15.49% | 1.52e-14% | 3.702% | 3.155% | 3.084% | 11.39% | 10.65%
| PDM3 |  | -2.184% | 8.032% | 2.163% | -2.184% | 1.863% | 41.29% | -19.57% | 1.076% | 6.663% | -23.92% | -14.57% | 1.402% | 13.03% | 8.032% | 8.684% | 30.42% | 13% | 1.62% | -5.445% | -26.09% | 24.99%
| Z-Score (Absolute) |  | NA | 0.67 | 0.06 | NA | 0.03 | NA | 2.25 | 0.06 | 0.54 | 2.71 | 1.72 | 0.03 | 1.21 | 0.67 | NA | NA | 1.21 | 0 | 0.75 | 2.94 | NA
Results from laboratories 1, 4, 6, 17, 18 and 26 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 553b (Indicative Value 0.33 ppm) | [23 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 03 | 06 | 08 | 09 | 17 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 50 | < 2 | 0.4 | < 10 | < 10 | 0.3 | < 1 | 1.6 | < 0.5 | [13 empty] | 1
| 1 | 2 | < 50 | < 2 | 0.3 | < 10 | < 10 | 0.3 | < 1 | 1.5 | < 0.5 | [13 empty] | 8
| 1 | 3 | < 50 | < 2 | 0.3 | < 10 | < 10 | 0.3 | < 1 | 1.3 | < 0.5 | [13 empty] | 16
| 1 | 4 | < 50 | < 2 | 0.3 | < 10 | < 10 | 0.3 | < 1 | 1.7 | < 0.5 | [13 empty] | 0.325
| 1 | 5 | < 50 | < 2 | 0.4 | < 10 | < 10 | 0.3 | < 1 | 1.2 | < 0.5 | [13 empty] | 27
| 1 | 6 | < 50 | < 2 | 0.3 | < 10 | < 10 | 0.4 | < 1 | 1.8 | < 0.5
| Mean |  |  |  | 0.33333 |  |  | 0.31667 |  | 1.5167
| Median |  |  |  | 0.3 |  |  | 0.3 |  | 1.55
| Std Dev. |  |  |  | 0.05164 |  |  | 0.040825 |  | 0.23166
| Rel.Std.Dev. |  |  |  | 15.49% |  |  | 12.89% |  | 15.27%
| PDM3 |  |  |  | 2.564% |  |  | -2.564% |  | 366.7%
| Z-Score (Absolute) |  | 23.68 | 0 | 0.66 | 3.95 | 3.95 | 0.67 | 0.49 | 0.51 | 0.74
| Analytical results for Ce in OREAS 553b (Certified Value 152 ppm) | [23 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 0
0
| 1 | 1 | 155 | 155 | 155 | 155.168 | 161 | 142 | 151.1 | 153 | 156 | [13 empty] | 1
| 1 | 2 | 151 | 147 | 157 | 154.128 | 155 | 145 | 151.3 | 161 | 145 | [13 empty] | 9
| 1 | 3 | 142 | 140 | 150 | 152.683 | 156 | 155 | 150.4 | 149 | 145 | [13 empty] | 16
| 1 | 4 | 148 | 124 | 151 | 153.4591 | 153 | 144 | 151.8 | 154 | 147 | [13 empty] | 151.6062
| 1 | 5 | 156 | 150 | 157 | 157.9455 | 157 | 149 | 151 | 149 | 153 | [13 empty] | 100
| 1 | 6 | 157 | 146 | 154 | 153.0512 | 160 | 143 | 152.1 | 154 | 148
| Mean |  | 151.5 | 143.6667 | 154 | 154.4058 | 157 | 146.3333 | 151.2833 | 153.3333 | 149
| Median |  | 153 | 146.5 | 154.5 | 153.7936 | 156.5 | 144.5 | 151.2 | 153.5 | 147.5
| Std Dev. |  | 5.7533 | 10.8197 | 2.9665 | 1.9435 | 3.0332 | 4.8854 | 0.6047 | 4.4121 | 4.5166
| Rel.Std.Dev. |  | 3.798% | 7.531% | 1.926% | 1.259% | 1.932% | 3.339% | 0.3997% | 2.877% | 3.031%
| PDM3 |  | -0.07005% | -5.237% | 1.579% | 1.847% | 3.558% | -3.478% | -0.213% | 1.139% | -1.719%
| Z-Score (Absolute) |  | 0 | 2.11 | 0.67 | 0.78 | 1.48 | 1.39 | 0.06 | 0.49 | 0.67
| Analytical results for Co in OREAS 553b (Certified Value 0.581 wt.%) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.626 | 0.533 | 0.58 | 0.578 | 0.55822 | 0.4078 | 0.602 | 0.582 | 0.60345 | 0.584 | 0.563 | 0.574 | 0.552 | 0.582 | 0.56938 | 0.59 | 0.5855 | 0.5966 | 0.599 | 0.53154 | 0.61 | 0.5656 | 1
| 1 | 2 | 0.6137 | 0.525 | 0.578 | 0.575 | 0.55834 | 0.4058 | 0.6027 | 0.589 | 0.59643 | 0.585 | 0.577 | 0.583 | 0.55 | 0.597 | 0.57221 | 0.59 | 0.6088 | 0.5972 | 0.597 | 0.54383 | 0.592 | 0.5634 | 30
| 1 | 3 | 0.6267 | 0.547 | 0.581 | 0.572 | 0.55959 | 0.4179 | 0.6023 | 0.583 | 0.59558 | 0.599 | 0.564 | 0.578 | 0.553 | 0.596 | 0.57054 | 0.59 | 0.5839 | 0.5996 | 0.595 | 0.53574 | 0.65 | 0.5581 | 16
| 1 | 4 | 0.6203 | 0.55 | 0.579 | 0.565 | 0.55781 | 0.412 | 0.6026 | 0.591 | 0.59791 | 0.595 | 0.573 | 0.574 | 0.557 | 0.581 | 0.57317 | 0.59 | 0.5854 | 0.6003 | 0.593 | 0.53691 | 0.618 | 0.5746 | 0.58122
| 1 | 5 | 0.6159 | 0.549 | 0.575 | 0.572 | 0.55727 | 0.4011 | 0.6036 | 0.582 | 0.60287 | 0.585 | 0.569 | 0.575 | 0.56 | 0.584 | 0.58249 | 0.59 | 0.5961 | 0.5914 | 0.608 | 0.52087 | 0.604 | 0.5765 | 101
| 1 | 6 | 0.642 | 0.561 | 0.576 | 0.566 | 0.55685 | 0.4135 | 0.6015 | 0.584 | 0.59546 | 0.593 | 0.574 | 0.586 | 0.556 | 0.591 | 0.57384 | 0.58 | 0.5895 | 0.5963 | 0.615 | 0.52174 | 0.506 | 0.5448
| Mean |  | 0.6241 | 0.54417 | 0.57817 | 0.57133 | 0.55801 | 0.40968 | 0.60245 | 0.58517 | 0.59861 | 0.59017 | 0.57 | 0.57833 | 0.55467 | 0.5885 | 0.57361 | 0.58833 | 0.59153 | 0.5969 | 0.60117 | 0.53177 | 0.59667 | 0.56383
| Median |  | 0.62315 | 0.548 | 0.5785 | 0.572 | 0.55801 | 0.4099 | 0.60245 | 0.5835 | 0.59717 | 0.589 | 0.571 | 0.5765 | 0.5545 | 0.5875 | 0.57269 | 0.59 | 0.5875 | 0.5969 | 0.598 | 0.53364 | 0.607 | 0.5645
| Std Dev. |  | 0.010204 | 0.012968 | 0.0023166 | 0.0050465 | 0.00095414 | 0.0059958 | 0.00071204 | 0.0038687 | 0.003632 | 0.0063377 | 0.0056569 | 0.0050859 | 0.0036697 | 0.0071204 | 0.0046558 | 0.0040825 | 0.0095437 | 0.0031509 | 0.0085421 | 0.0090232 | 0.048525 | 0.011617
| Rel.Std.Dev. |  | 1.635% | 2.383% | 0.4007% | 0.8833% | 0.171% | 1.464% | 0.1182% | 0.6611% | 0.6067% | 1.074% | 0.9924% | 0.8794% | 0.6616% | 1.21% | 0.8117% | 0.6939% | 1.613% | 0.5279% | 1.421% | 1.697% | 8.133% | 2.06%
| PDM3 |  | 7.378% | -6.375% | -0.5253% | -1.701% | -3.993% | -29.51% | 3.653% | 0.679% | 2.993% | 1.539% | -1.93% | -0.4967% | -4.569% | 1.253% | -1.31% | 1.224% | 1.774% | 2.698% | 3.432% | -8.508% | 2.658% | -2.991%
| Z-Score (Absolute) |  | 1.9 | 1.69 | 0.16 | 0.47 | 1.06 | 7.72 | 0.93 | 0.15 | 0.76 | 0.38 | 0.53 | 0.15 | 1.21 | 0.3 | 0.37 | 0.3 | 0.44 | 0.68 | 0.87 | 2.24 | 0.67 | 0.8
| Analytical results for Cr in OREAS 553b (Certified Value 110 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 25 | 26 |  |  |  | ppm
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  |  |  | 0
0
| 1 | 1 | 112 | 116 | 100 | 90 | 155 | 102 | 90 | 95.6369 | 140 | 210 | 140 | < 70 | 210 | 106 | 110 | 99 | < 100 | 119 | 108 |  |  |  | 1
| 1 | 2 | 111 | 116 | 120 | 90 | 115 | 102 | 100 | 92.8167 | 140 | 140 | 140 | < 70 | 140 | 106 | 120 | 112 | < 100 | 109 | 110 |  |  |  | #N/A
| 1 | 3 | 106 | 116 | 100 | 90 | 130 | 102 | 90 | 98.1651 | 140 | 140 | 140 | < 70 | 140 | 105 | 108 | 108 | < 100 | 136 | 120 |  |  |  | 16
| 1 | 4 | 104 | 116 | 120 | 90 | 130 | 102 | 100 | 91.8151 | 140 | 140 | 140 | < 70 | 140 | 108 | 104 | 102 | < 100 | 113 | 116 |  |  |  | 110.467
| 1 | 5 | 111 | 116 | 120 | 90 | 115 | 102 | 90 | 96.9209 | 140 | 210 | 140 | 70 | 140 | 107 | 108 | 108 | < 100 | 133 | 109 |  |  |  | 102
| 1 | 6 | 109 | 116 | 100 | 90 | 124 | 102 | 90 | 94.6084 | 140 | 210 | 140 | < 70 | 140 | 105 | 113 | 102 | < 100 | 124 | 104
| Mean |  | 108.8333 | 116 | 110 | 90 | 128.1667 | 102 | 93.3333 | 94.9939 | 140 | 175 | 140 | 70 | 151.6667 | 106.1667 | 110.5 | 105.1667 |  | 122.3333 | 111.1667
| Median |  | 110 | 116 | 110 | 90 | 127 | 102 | 90 | 95.1227 | 140 | 175 | 140 | 70 | 140 | 106 | 109 | 105 |  | 121.5 | 109.5
| Std Dev. |  | 3.1885 | 0 | 10.9545 | 0 | 14.7705 | 0 | 5.164 | 2.4157 | 0 | 38.3406 | 0 |  | 28.5774 | 1.169 | 5.5045 | 4.916 |  | 10.7641 | 5.8109
| Rel.Std.Dev. |  | 2.93% | 0% | 9.959% | 0% | 11.52% | 0% | 5.533% | 2.543% | 0% | 21.91% | 0% |  | 18.84% | 1.101% | 4.981% | 4.674% |  | 8.799% | 5.227%
| PDM3 |  | -1.479% | 5.009% | -0.4228% | -18.53% | 16.02% | -7.665% | -15.51% | -14.01% | 26.73% | 58.42% | 26.73% | -36.63% | 37.3% | -3.893% | 0.02985% | -4.798% |  | 10.74% | 0.6334%
| Z-Score (Absolute) |  | 0.05 | 0.28 | 0 | 0.9 | 0.82 | 0.36 | 0.75 | 0.67 | 1.21 | 2.74 | 1.21 | 3.15 | 1.72 | 0.17 | 0.02 | 0.22 | 2.7 | 0.55 | 0.05
| Analytical results for Cs in OREAS 553b (Certified Value 2.13 ppm) | [23 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 14 | 15 | 17 | 18 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 2
3
| 1 | 1 | 2.4 | 2.2 | 2.13 | 2.1 | 1.8432 | 2.4 | < 5 | 2.0358 | 2.1 | 2 | 2.1 | 2.3 | [10 empty] | 1
| 1 | 2 | 2.2 | 2.1 | 2.15 | 2.1 | 1.918 | 2.1 | < 5 | 2.0399 | 2.1 | 2.4 | 2.1 | 2.1 | [10 empty] | 12
| 1 | 3 | 2.2 | 2.9 | 2.19 | 2.2 | 1.8359 | 2.2 | < 5 | 2.0644 | 2.1 | 2.5 | 2.2 | 2.4 | [10 empty] | 16
| 1 | 4 | 2.1 | 2.1 | 2.11 | 2.1 | 1.8203 | 2 | < 5 | 2.0355 | 2 | 2.4 | 2.1 | 2.1 | [10 empty] | 2.1337
| 1 | 5 | 2.3 | 2.6 | 2.11 | 2.2 | 1.9588 | 2.4 | < 5 | 2.0646 | 2.1 | 2.2 | 2.1 | 2.2 | [10 empty] | 103
| 1 | 6 | 2.2 | 2.1 | 2.15 | 2.1 | 1.9059 | 2.2 | < 5 | 1.9918 | 2.1 | 2.3 | 1.7 | 2.1
| Mean |  | 2.2333 | 2.3333 | 2.14 | 2.1333 | 1.8803 | 2.2167 |  | 2.0387 | 2.0833 | 2.3 | 2.05 | 2.2
| Median |  | 2.2 | 2.15 | 2.14 | 2.1 | 1.8746 | 2.2 |  | 2.0378 | 2.1 | 2.35 | 2.1 | 2.15
| Std Dev. |  | 0.10328 | 0.33862 | 0.030332 | 0.05164 | 0.055098 | 0.16021 |  | 0.026656 | 0.040825 | 0.17889 | 0.17607 | 0.12649
| Rel.Std.Dev. |  | 4.624% | 14.51% | 1.417% | 2.421% | 2.93% | 7.227% |  | 1.308% | 1.96% | 7.778% | 8.589% | 5.75%
| PDM3 |  | 4.67% | 9.356% | 0.2954% | -0.01706% | -11.87% | 3.889% |  | -4.454% | -2.36% | 7.794% | -3.923% | 3.107%
| Z-Score (Absolute) |  | 0.41 | 1.07 | 0.2 | 0.24 | 1.89 | 0.3 | 2.15 | 0.86 | 0.57 | 0.85 | 0.78 | 0.2
| Analytical results for Cu in OREAS 553b (Certified Value 3.61 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 3.6946 | 3.479 | 3.7 | 3.72 | 3.5312 | 4.2665 | 3.65 | 3.686 | 3.6828 | 3.54 | 3.48 | 3.49 | 3.4 | 3.64 | 3.63 | 3.7068 | 3.212 | 3.493 | 3.9153 | 3.3103 | 3.8715 |  | 1
| 1 | 2 | 3.7114 | 3.434 | 3.64 | 3.66 | 3.5311 | 4.1311 | 3.65 | 3.765 | 3.5523 | 3.48 | 3.47 | 3.64 | 3.37 | 3.68 | 3.59 | 3.7549 | 3.2691 | 3.477 | 3.8954 | 3.3187 | 3.8516 |  | 31
| 1 | 3 | 3.6312 | 3.65 | 3.62 | 3.6 | 3.5267 | 4.3746 | 3.65 | 3.857 | 3.7256 | 3.62 | 3.47 | 3.55 | 3.36 | 3.66 | 3.59 | 3.6202 | 3.4886 | 3.488 | 3.9823 | 3.3498 | 3.8737 |  | 16
| 1 | 4 | 3.6058 | 3.582 | 3.62 | 3.7 | 3.5334 | 4.0509 | 3.66 | 3.739 | 3.6102 | 3.53 | 3.53 | 3.53 | 3.39 | 3.56 | 3.61 | 3.5981 | 3.265 | 3.463 | 3.9012 | 3.3508 | 3.8266 |  | 3.6061
| 1 | 5 | 3.6552 | 3.564 | 3.72 | 3.65 | 3.5357 | 3.8919 | 3.66 | 3.73 | 3.6096 | 3.51 | 3.44 | 3.55 | 3.41 | 3.59 | 3.58 | 3.6774 | 3.4096 | 3.522 | 3.9124 | 3.3588 | 3.7569 |  | 104
| 1 | 6 | 3.6771 | 3.61 | 3.69 | 3.64 | 3.5404 | 4.3966 | 3.67 | 3.673 | 3.6948 | 3.62 | 3.47 | 3.57 | 3.38 | 3.64 | 3.62 | 3.6352 | 3.2594 | 3.554 | 3.9156 | 3.3054 | 3.7432
| Mean |  | 3.6625 | 3.5532 | 3.665 | 3.6617 | 3.5331 | 4.1853 | 3.6567 | 3.7417 | 3.6459 | 3.55 | 3.4767 | 3.555 | 3.385 | 3.6283 | 3.6033 | 3.6654 | 3.3173 | 3.4995 | 3.9204 | 3.3323 | 3.8206
| Median |  | 3.6662 | 3.573 | 3.665 | 3.655 | 3.5323 | 4.1988 | 3.655 | 3.7345 | 3.6465 | 3.535 | 3.47 | 3.55 | 3.385 | 3.64 | 3.6 | 3.6563 | 3.267 | 3.4905 | 3.9139 | 3.3343 | 3.8391
| Std Dev. |  | 0.039706 | 0.081556 | 0.043704 | 0.043089 | 0.0046776 | 0.19706 | 0.008165 | 0.066035 | 0.065499 | 0.057966 | 0.029439 | 0.049699 | 0.018708 | 0.044907 | 0.019664 | 0.058911 | 0.10711 | 0.033135 | 0.031427 | 0.02342 | 0.057369
| Rel.Std.Dev. |  | 1.084% | 2.295% | 1.192% | 1.177% | 0.1324% | 4.708% | 0.2233% | 1.765% | 1.797% | 1.633% | 0.8468% | 1.398% | 0.5527% | 1.238% | 0.5457% | 1.607% | 3.229% | 0.9468% | 0.8016% | 0.7028% | 1.502%
| PDM3 |  | 1.566% | -1.468% | 1.634% | 1.541% | -2.024% | 16.06% | 1.403% | 3.76% | 1.103% | -1.555% | -3.589% | -1.417% | -6.131% | 0.6168% | -0.07647% | 1.646% | -8.009% | -2.956% | 8.715% | -7.592% | 5.948%
| Z-Score (Absolute) |  | 0.29 | 0.65 | 0.32 | 0.29 | 0.82 | 4.79 | 0.24 | 0.98 | 0.15 | 0.67 | 1.31 | 0.63 | 2.09 | 0 | 0.22 | 0.32 | 2.68 | 1.11 | 2.51 | 2.55 | 1.65
| Analytical results for Dy in OREAS 553b (Certified Value 7.03 ppm) | [23 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 6.6 | 7.05 | 6.9956 | 7.7 | 7 | 5 | 7.18 | 6.85 | [14 empty] | 1
| 1 | 2 | 7.1 | 6.88 | 6.8231 | 7.29 | 7 | 5.2 | 7.7 | 6.74 | [14 empty] | 13
| 1 | 3 | 7.6 | 7.06 | 7.1491 | 7.03 | 7 | 5.3 | 6.88 | 6.79 | [14 empty] | 16
| 1 | 4 | 6.6 | 7.1 | 7.514 | 7.02 | 7 | 5.3 | 7.16 | 6.49 | [14 empty] | 7.0307
| 1 | 5 | 7.4 | 6.95 | 7.0573 | 7.33 | 7 | 5.2 | 6.97 | 7.15 | [14 empty] | 105
| 1 | 6 | 6.4 | 7.16 | 6.8263 | 7.43 | 7 | 5.2 | 7.16 | 6.6
| Mean |  | 6.95 | 7.0333 | 7.0609 | 7.3 | 7 | 5.2 | 7.175 | 6.77
| Median |  | 6.85 | 7.055 | 7.0265 | 7.31 | 7 | 5.2 | 7.16 | 6.765
| Std Dev. |  | 0.48888 | 0.10191 | 0.25651 | 0.25659 | 0 | 0.10954 | 0.28466 | 0.22777
| Rel.Std.Dev. |  | 7.034% | 1.449% | 3.633% | 3.515% | 0% | 2.107% | 3.967% | 3.364%
| PDM3 |  | -1.148% | 0.03737% | 0.4295% | 3.83% | -0.4367% | -26.04% | 2.052% | -3.708%
| Z-Score (Absolute) |  | 0.4 | 0 | 0.13 | 1.27 | NA | 8.73 | 0.67 | 1.25
| Analytical results for Er in OREAS 553b (Certified Value 3.73 ppm) | [23 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 3.9 | 3.79 | 3.9063 | 4.35 | 4 | 3.3 | 3.98 | 3.54 | [14 empty] | 1
| 1 | 2 | 4.9 | 3.49 | 3.8685 | 4.1 | 4 | 3.1 | 4.18 | 3.52 | [14 empty] | 14
| 1 | 3 | 3.8 | 3.79 | 3.8947 | 4.03 | 4 | 3.2 | 3.83 | 3.5 | [14 empty] | 16
| 1 | 4 | 3.4 | 3.52 | 3.8689 | 3.98 | 3 | 3.2 | 4.05 | 3.52 | [14 empty] | 3.7332
| 1 | 5 | 4.3 | 3.58 | 3.8562 | 4.2 | 4 | 3.2 | 3.72 | 3.71 | [14 empty] | 106
| 1 | 6 | 3.3 | 3.77 | 3.8613 | 4.26 | 4 | 3.3 | 3.9 | 3.49
| Mean |  | 3.9333 | 3.6567 | 3.876 | 4.1533 | 3.8333 | 3.2167 | 3.9433 | 3.5467
| Median |  | 3.85 | 3.675 | 3.8687 | 4.15 | 4 | 3.2 | 3.94 | 3.52
| Std Dev. |  | 0.59554 | 0.14194 | 0.019916 | 0.14166 | 0.40825 | 0.075277 | 0.1633 | 0.081894
| Rel.Std.Dev. |  | 15.14% | 3.882% | 0.5138% | 3.411% | 10.65% | 2.34% | 4.141% | 2.309%
| PDM3 |  | 5.36% | -2.051% | 3.823% | 11.25% | 2.681% | -13.84% | 5.628% | -4.997%
| Z-Score (Absolute) |  | 0.18 | 0.67 | 0 | 0.85 | NA | 2.03 | 0.21 | 1.01
| Analytical results for Eu in OREAS 553b (Certified Value 0.91 ppm) | [23 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
2
| 1 | 1 | 1 | 0.75 | 0.93773 | 1.06 | < 2 | 0.9 | 0.86 | 0.96 | [14 empty] | 1
| 1 | 2 | 0.9 | 0.79 | 0.93057 | 1.01 | < 2 | 0.9 | 0.91 | 0.86 | [14 empty] | 15
| 1 | 3 | 0.9 | 0.79 | 0.98981 | 1.04 | < 2 | 0.9 | 0.93 | 0.85 | [14 empty] | 16
| 1 | 4 | 0.8 | 0.74 | 0.97491 | 1 | < 2 | 0.9 | 1.06 | 0.87 | [14 empty] | 0.91026
| 1 | 5 | 1.1 | 0.8 | 0.95211 | 1.02 | < 2 | 0.9 | 0.92 | 0.89 | [14 empty] | 107
| 1 | 6 | 0.7 | 0.76 | 0.89566 | 1.07 | < 2 | 0.9 | 0.94 | 0.87
| Mean |  | 0.9 | 0.77167 | 0.9468 | 1.0333 |  | 0.9 | 0.93667 | 0.88333
| Median |  | 0.9 | 0.775 | 0.94492 | 1.03 |  | 0.9 | 0.925 | 0.87
| Std Dev. |  | 0.14142 | 0.024833 | 0.033553 | 0.028048 |  | 0 | 0.066533 | 0.039833
| Rel.Std.Dev. |  | 15.71% | 3.218% | 3.544% | 2.714% |  | 0% | 7.103% | 4.509%
| PDM3 |  | -1.127% | -15.23% | 4.014% | 13.52% |  | -1.127% | 2.901% | -2.958%
| Z-Score (Absolute) |  | 0.39 | 3.12 | 0.6 | 2.44 | 1.74 | 0.39 | 0.39 | 0.74
| Analytical results for Fe in OREAS 553b (Certified Value 3.17 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 3.28 | 3 | 3.26 | 3.28 | 3.1211 | 3.7 | 3.14 | 3.12 | 3.1606 | 3.078 | 3.022 | 3.141 | 3.12 | 3.099 | 3.2891 | 3.17 | 3.22 | 3.0587 | 3.04 | 2.81 | 3.27 | 3.32 | 1
| 1 | 2 | 3.32 | 2.95 | 3.25 | 3.3 | 3.1276 | 3.17 | 3.13 | 3.22 | 3.0904 | 3.12 | 3.029 | 3.05 | 3.036 | 3.217 | 3.2539 | 3.15 | 3.31 | 3.0931 | 3.03 | 2.82 | 3.27 | 3.35 | #N/A
| 1 | 3 | 3.23 | 3.03 | 3.24 | 3.23 | 3.119 | 3.6 | 3.13 | 3.16 | 3.3106 | 3.189 | 2.973 | 3.12 | 3.071 | 3.21 | 3.2308 | 3.14 | 3.21 | 3.305 | 3.05 | 2.8 | 3.4 | 3.37 | 16
| 1 | 4 | 3.23 | 3.06 | 3.25 | 3.2 | 3.1194 | 3.43 | 3.13 | 3.23 | 3.152 | 3.189 | 3.008 | 3.127 | 3.113 | 3.099 | 3.1671 | 3.16 | 3.13 | 3.0766 | 3.02 | 2.83 | 3.23 | 3.32 | 3.1677
| 1 | 5 | 3.27 | 3.04 | 3.26 | 3.23 | 3.1228 | 3.51 | 3.14 | 3.14 | 3.216 | 3.127 | 2.98 | 3.217 | 3.099 | 3.182 | 3.2377 | 3.15 | 3.19 | 3.2281 | 3.08 | 2.81 | 3.27 | 3.33 | 108
| 1 | 6 | 3.27 | 3.13 | 3.22 | 3.22 | 3.1286 | 3.39 | 3.12 | 3.25 | 3.2542 | 3.148 | 3.057 | 3.231 | 3.092 | 3.217 | 3.2622 | 3.18 | 3.18 | 3.0558 | 3.1 | 8.82 | 3.22 | 3.22
| Mean |  | 3.2667 | 3.035 | 3.2467 | 3.2433 | 3.1231 | 3.4667 | 3.1317 | 3.1867 | 3.1973 | 3.1418 | 3.0115 | 3.1477 | 3.0885 | 3.1707 | 3.2401 | 3.1583 | 3.2067 | 3.1362 | 3.0533 | 3.815 | 3.2767 | 3.3183
| Median |  | 3.27 | 3.035 | 3.25 | 3.23 | 3.1219 | 3.47 | 3.13 | 3.19 | 3.1883 | 3.1375 | 3.015 | 3.134 | 3.0955 | 3.196 | 3.2458 | 3.155 | 3.2 | 3.0849 | 3.045 | 2.815 | 3.27 | 3.325
| Std Dev. |  | 0.033862 | 0.060249 | 0.015055 | 0.038297 | 0.004132 | 0.18403 | 0.0075277 | 0.053541 | 0.079062 | 0.043032 | 0.031539 | 0.067123 | 0.030911 | 0.056994 | 0.041224 | 0.01472 | 0.059554 | 0.10471 | 0.030768 | 2.452 | 0.064395 | 0.051929
| Rel.Std.Dev. |  | 1.037% | 1.985% | 0.4637% | 1.181% | 0.1323% | 5.309% | 0.2404% | 1.68% | 2.473% | 1.37% | 1.047% | 2.132% | 1.001% | 1.798% | 1.272% | 0.4661% | 1.857% | 3.339% | 1.008% | 64.27% | 1.965% | 1.565%
| PDM3 |  | 3.123% | -4.19% | 2.492% | 2.387% | -1.41% | 9.437% | -1.138% | 0.5978% | 0.9333% | -0.8175% | -4.932% | -0.6333% | -2.501% | 0.09275% | 2.285% | -0.2966% | 1.229% | -0.9948% | -3.611% | 20.43% | 3.439% | 4.754%
| Z-Score (Absolute) |  | 0.94 | 1.54 | 0.73 | 0.69 | 0.6 | 3.08 | 0.5 | 0.08 | 0.2 | 0.4 | 1.79 | 0.33 | 0.97 | 0.08 | 0.66 | 0.22 | 0.3 | 0.46 | 1.34 | 6.81 | 1.05 | 1.49
| Analytical results for Ga in OREAS 553b (Certified Value 27.7 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 1
2
| 1 | 1 | 28 | 30.8 | 28 | 24 | 20 | 28.2032 | 27.1058 | 27 | 28 | 29.5 | 27 | 27 | [10 empty] | 1
| 1 | 2 | 27 | 26.3 | 27 | 24 | 19 | 27.0352 | 26.956 | 27 | 27 | 29.5 | 27 | 26 | [10 empty] | 17
| 1 | 3 | 28 | 34.5 | 28 | 25 | 20 | 28.6322 | 26.6154 | 27 | 27 | 29.2 | 29 | 26 | [10 empty] | 16
| 1 | 4 | 28 | 27.7 | 28 | 25 | 19 | 26.9343 | 26.6183 | 28 | 26 | 30.2 | 27 | 26 | [10 empty] | 27.6572
| 1 | 5 | 28 | 34 | 28 | 25 | 20 | 29.1539 | 26.4075 | 28 | 28 | 29.8 | 27 | 27 | [10 empty] | 109
| 1 | 6 | 29 | 28 | 28 | 25 | 20 | 28.3987 | 26.7293 | 28 | 28 | 30.1 | 22 | 26
| Mean |  | 28 | 30.2167 | 27.8333 | 24.6667 | 19.6667 | 28.0596 | 26.7387 | 27.5 | 27.3333 | 29.7167 | 26.5 | 26.3333
| Median |  | 28 | 29.4 | 28 | 25 | 20 | 28.3009 | 26.6738 | 27.5 | 27.5 | 29.65 | 27 | 26
| Std Dev. |  | 0.63246 | 3.4522 | 0.40825 | 0.5164 | 0.5164 | 0.89181 | 0.25358 | 0.54772 | 0.8165 | 0.38687 | 2.3452 | 0.5164
| Rel.Std.Dev. |  | 2.259% | 11.42% | 1.467% | 2.094% | 2.626% | 3.178% | 0.9484% | 1.992% | 2.987% | 1.302% | 8.85% | 1.961%
| PDM3 |  | 1.239% | 9.254% | 0.6368% | -10.81% | -28.89% | 1.455% | -3.321% | -0.5685% | -1.171% | 7.446% | -4.184% | -4.787%
| Z-Score (Absolute) |  | 0.49 | 2.37 | 0.35 | 2.33 | 6.55 | 0.54 | 0.57 | 0.07 | 0.07 | 1.95 | 0.78 | 0.92
| Analytical results for Gd in OREAS 553b (Certified Value 6.87 ppm) | [23 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 22 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [15 empty] | 2
3
| 1 | 1 | 7.9 | 6.22 | 7.1123 | 7.17 | 5.6 | 6.97 | 6.84 | [15 empty] | 1
| 1 | 2 | 7.6 | 6.21 | 7.2316 | 6.86 | 5.7 | 7.15 | 6.4 | [15 empty] | 18
| 1 | 3 | 6.9 | 6.49 | 7.2602 | 6.82 | 5.9 | 6.8 | 6.6 | [15 empty] | 16
| 1 | 4 | 6.2 | 6.41 | 7.6024 | 6.56 | 5.8 | 6.7 | 6.64 | [15 empty] | 6.8708
| 1 | 5 | 7.4 | 6.47 | 7.4084 | 7.01 | 5.8 | 6.93 | 6.75 | [15 empty] | 110
| 1 | 6 | 6.5 | 6.38 | 7.1456 | 7.09 | 6 | 6.99 | 6.63
| Mean |  | 7.0833 | 6.3633 | 7.2934 | 6.9183 | 5.8 | 6.9233 | 6.6433
| Median |  | 7.15 | 6.395 | 7.2459 | 6.935 | 5.8 | 6.95 | 6.635
| Std Dev. |  | 0.66156 | 0.1216 | 0.18347 | 0.22031 | 0.14142 | 0.15693 | 0.14922
| Rel.Std.Dev. |  | 9.34% | 1.911% | 2.516% | 3.184% | 2.438% | 2.267% | 2.246%
| PDM3 |  | 3.093% | -7.386% | 6.15% | 0.6911% | -15.59% | 0.7639% | -3.311%
| Z-Score (Absolute) |  | 0.4 | 1.36 | 0.92 | 0 | 2.74 | 0.01 | 0.67
| Analytical results for Ge in OREAS 553b (Certified Value 2.24 ppm) | [23 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 17 | 18 | 25 | 26 | [14 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
2
| 1 | 1 | 2 | 2.2 | 2 | 1.8145 | 2 | 2 | < 5 | 4 | [14 empty] | 1
| 1 | 2 | 2 | 3.4 | 2 | 1.7087 | 2 | 2 | < 5 | 3 | [14 empty] | 2
| 1 | 3 | 2 | 1.8 | 2 | 1.8343 | 2 | 2 | < 5 | 4 | [14 empty] | 16
| 1 | 4 | 2 | 2.6 | 2 | 1.7305 | 2 | 2 | < 5 | 3 | [14 empty] | 2.2382
| 1 | 5 | 2 | 1.9 | 2 | 1.7023 | 2 | 2 | < 5 | 4 | [14 empty] | 111
| 1 | 6 | 2 | 1.5 | 2 | 1.8129 | 2 | 2 | < 5 | 4
| Mean |  | 2 | 2.2333 | 2 | 1.7672 | 2 | 2 |  | 3.6667
| Median |  | 2 | 2.05 | 2 | 1.7717 | 2 | 2 |  | 4
| Std Dev. |  | 0 | 0.68313 | 0 | 0.059675 | 0 | 0 |  | 0.5164
| Rel.Std.Dev. |  | 0% | 30.59% | 0% | 3.377% | 0% | 0% |  | 14.08%
| PDM3 |  | -10.64% | -0.2161% | -10.64% | -21.04% | -10.64% | -10.64% |  | 63.82%
| Z-Score (Absolute) |  | 0 | 1.35 | 0 | 1.35 | 0 | 0 | 2.9 | 9.65
| Analytical results for Hf in OREAS 553b (Indicative Value 7.97 ppm) | [23 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 03 | 15 | 17 | 22 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | [18 empty] | 2
2
| 1 | 1 | 10 | 7.157 | 8 | 9 | [18 empty] | 1
| 1 | 2 | 10 | 6.6754 | 7 | 8.7 | [18 empty] | 22
| 1 | 3 | 10 | 6.6879 | 7 | 8.8 | [18 empty] | 16
| 1 | 4 | 20 | 7.1672 | 7 | 8.9 | [18 empty] | 7.9715
| 1 | 5 | 10 | 6.8361 | 7 | 8.9 | [18 empty] | 28
| 1 | 6 | < 10 | 6.8913 | 7 | 8.6
| Mean |  | 12 | 6.9025 | 7.1667 | 8.8167
| Median |  | 10 | 6.8637 | 7 | 8.85
| Std Dev. |  | 4.4721 | 0.21771 | 0.40825 | 0.1472
| Rel.Std.Dev. |  | 37.27% | 3.154% | 5.696% | 1.67%
| PDM3 |  | 50.54% | -13.41% | -10.1% | 10.6%
| Z-Score (Absolute) |  | 2 | 0.77 | 0.58 | 0.58
| Analytical results for Hg in OREAS 553b (Indicative Value < 5 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | Lab | [21 empty] | 1
| No | No | 22 | [21 empty] | ppm
|  |  | PF*OES | [21 empty] | 2
2
| 1 | 1 | < 5 | [21 empty] | 1
| 1 | 2 | < 5 | [21 empty] | 23
| 1 | 3 | < 5 | [21 empty] | 16
| 1 | 4 | < 5 | [21 empty] | < 5
| 1 | 5 | < 5 | [21 empty] | 29
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 553b (Certified Value 1.35 ppm) | [23 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 22 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [15 empty] | 2
3
| 1 | 1 | 1.5 | 1.24 | 1.4123 | 1.56 | 1.1 | 1.43 | 1.3 | [15 empty] | 1
| 1 | 2 | 1.5 | 1.23 | 1.3384 | 1.49 | 1 | 1.37 | 1.28 | [15 empty] | 5
| 1 | 3 | 1.4 | 1.33 | 1.2788 | 1.42 | 1 | 1.33 | 1.27 | [15 empty] | 16
| 1 | 4 | 1.1 | 1.25 | 1.4466 | 1.45 | 1.1 | 1.4 | 1.29 | [15 empty] | 1.352
| 1 | 5 | 1.4 | 1.23 | 1.3706 | 1.5 | 1 | 1.31 | 1.34 | [15 empty] | 112
| 1 | 6 | 1.1 | 1.29 | 1.3856 | 1.48 | 1.1 | 1.38 | 1.27
| Mean |  | 1.3333 | 1.2617 | 1.3721 | 1.4833 | 1.05 | 1.37 | 1.2917
| Median |  | 1.4 | 1.245 | 1.3781 | 1.485 | 1.05 | 1.375 | 1.285
| Std Dev. |  | 0.18619 | 0.040208 | 0.058626 | 0.04761 | 0.054772 | 0.044272 | 0.026394
| Rel.Std.Dev. |  | 13.96% | 3.187% | 4.273% | 3.21% | 5.216% | 3.232% | 2.043%
| PDM3 |  | -1.381% | -6.682% | 1.482% | 9.713% | -22.34% | 1.331% | -4.463%
| Z-Score (Absolute) |  | 0 | 1.16 | 0.63 | 2.43 | 4.59 | 0.59 | 0.67
| Analytical results for In in OREAS 553b (Certified Value 2.22 ppm) | [23 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 17 | 18 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 2.3 | 2.4 | 2.5 | 1.8 | 1.8902 | 2.1 | 2.1 | 2.43 | 2.1 | [13 empty] | 1
| 1 | 2 | 2.2 | 2.1 | 2.4 | 1.8 | 1.9635 | 2.1 | 2.2 | 2.28 | 2.2 | [13 empty] | 6
| 1 | 3 | 2.1 | 2.5 | 2.4 | 1.8 | 1.8826 | 2.1 | 2.2 | 2.41 | 2.2 | [13 empty] | 16
| 1 | 4 | 2.1 | 2.5 | 2.4 | 1.8 | 1.9694 | 2.1 | 2.1 | 2.21 | 2.2 | [13 empty] | 2.2177
| 1 | 5 | 2.2 | 2.2 | 2.5 | 1.6 | 2.011 | 2.1 | 2.3 | 2.33 | 2.4 | [13 empty] | 113
| 1 | 6 | 2.2 | 2.1 | 2.5 | 1.7 | 1.9825 | 2.2 | 2.3 | 2.29 | 2.2
| Mean |  | 2.1833 | 2.3 | 2.45 | 1.75 | 1.9499 | 2.1167 | 2.2 | 2.325 | 2.2167
| Median |  | 2.2 | 2.3 | 2.45 | 1.8 | 1.9664 | 2.1 | 2.2 | 2.31 | 2.2
| Std Dev. |  | 0.075277 | 0.18974 | 0.054772 | 0.083666 | 0.051871 | 0.040825 | 0.089443 | 0.083367 | 0.098319
| Rel.Std.Dev. |  | 3.448% | 8.249% | 2.236% | 4.781% | 2.66% | 1.929% | 4.066% | 3.586% | 4.435%
| PDM3 |  | -1.549% | 3.711% | 10.48% | -21.09% | -12.08% | -4.555% | -0.7978% | 4.839% | -0.04626%
| Z-Score (Absolute) |  | 0.11 | 0.67 | 1.69 | 3.03 | 1.69 | 0.56 | 0 | 0.84 | 0.11
| Analytical results for K in OREAS 553b (Certified Value 2.58 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 2.58 | 2.656 | 2.6 | 2.72 | 3.1 | 2.63 | 2.49 | 2.4097 | 2.582 | 2.59 | 2.632 | 2.407 | 2.706 | 2.6 | 2.63 | 2.2172 | 2.61 | 0.21 | 2.34 | 2.7 |  |  | 1
| 1 | 2 | 2.59 | 2.557 | 2.6 | 2.73 | 2.5 | 2.61 | 2.45 | 2.4608 | 2.573 | 2.598 | 2.549 | 2.432 | 2.656 | 2.6 | 2.71 | 2.3324 | 2.6 | 0.22 | 2.34 | 2.8 |  |  | #N/A
| 1 | 3 | 2.55 | 2.507 | 2.6 | 2.7 | 2.7 | 2.62 | 2.52 | 2.4352 | 2.64 | 2.549 | 2.532 | 2.407 | 2.69 | 2.6 | 2.63 | 2.4752 | 2.61 | 0.21 | 2.46 | 2.8 |  |  | 16
| 1 | 4 | 2.56 | 2.549 | 2.6 | 2.71 | 2.6 | 2.61 | 2.48 | 2.3496 | 2.59 | 2.573 | 2.573 | 2.407 | 2.607 | 2.6 | 2.57 | 2.2883 | 2.59 | 0.21 | 2.32 | 2.7 |  |  | 2.5832
| 1 | 5 | 2.59 | 2.549 | 2.6 | 2.62 | 2.4 | 2.62 | 2.45 | 2.3669 | 2.598 | 2.549 | 2.582 | 2.399 | 2.632 | 2.6 | 2.62 | 2.4122 | 2.66 | 0.21 | 2.31 | 2.7 |  |  | 114
| 1 | 6 | 2.55 | 2.573 | 2.6 | 2.64 | 2.8 | 2.61 | 2.51 | 2.4054 | 2.64 | 2.59 | 2.598 | 2.383 | 2.656 | 2.6 | 2.6 | 2.2796 | 2.68 | 0.21 | 2.3 | 2.6
| Mean |  | 2.57 | 2.5652 | 2.6 | 2.6867 | 2.6833 | 2.6167 | 2.4833 | 2.4046 | 2.6038 | 2.5748 | 2.5777 | 2.4058 | 2.6578 | 2.6 | 2.6267 | 2.3342 | 2.625 | 0.21167 | 2.345 | 2.7167
| Median |  | 2.57 | 2.553 | 2.6 | 2.705 | 2.65 | 2.615 | 2.485 | 2.4076 | 2.594 | 2.5815 | 2.5775 | 2.407 | 2.656 | 2.6 | 2.625 | 2.3104 | 2.61 | 0.21 | 2.33 | 2.7
| Std Dev. |  | 0.018974 | 0.049568 | 0 | 0.045461 | 0.24833 | 0.008165 | 0.029439 | 0.041378 | 0.029219 | 0.021609 | 0.035534 | 0.015855 | 0.036367 | 0 | 0.046762 | 0.094599 | 0.036194 | 0.0040825 | 0.058566 | 0.075277
| Rel.Std.Dev. |  | 0.7383% | 1.932% | 0% | 1.692% | 9.254% | 0.312% | 1.185% | 1.721% | 1.122% | 0.8393% | 1.379% | 0.659% | 1.368% | 0% | 1.78% | 4.053% | 1.379% | 1.929% | 2.497% | 2.771%
| PDM3 |  | -0.5125% | -0.6996% | 0.6488% | 4.004% | 3.875% | 1.294% | -3.867% | -6.915% | 0.7972% | -0.3254% | -0.2157% | -6.868% | 2.888% | 0.6488% | 1.681% | -9.643% | 1.617% | -91.81% | -9.223% | 5.165%
| Z-Score (Absolute) |  | 0.24 | 0.3 | 0.14 | 1.23 | 1.19 | 0.35 | 1.33 | 2.32 | 0.19 | 0.18 | 0.14 | 2.31 | 0.87 | 0.14 | 0.48 | 3.21 | 0.46 | 29.99 | 3.08 | 1.61
| Analytical results for La in OREAS 553b (Certified Value 77 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 15 | 17 | 19 | 22 | 25 | 26 | [11 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [11 empty] | 0
1
| 1 | 1 | 75 | 84.2 | 82.8 | 74 | 70 | 75.8204 | 84.5 | 71 | 76.4 | 76.9 | 80.8 | [11 empty] | 1
| 1 | 2 | 72 | 79.8 | 80 | 74 | 70 | 75.654 | 79.1 | 72 | 77.2 | 81.3 | 81.3 | [11 empty] | 7
| 1 | 3 | 71 | 83 | 81 | 74 | 70 | 74.2565 | 80.8 | 77 | 76 | 75.4 | 83.4 | [11 empty] | 16
| 1 | 4 | 73 | 78.2 | 80.7 | 74 | 80 | 77.0997 | 80.8 | 72 | 77.9 | 76 | 82.4 | [11 empty] | 76.9894
| 1 | 5 | 73 | 82 | 81.2 | 74 | 80 | 73.5764 | 81.8 | 75 | 77.5 | 75 | 87.1 | [11 empty] | 115
| 1 | 6 | 73 | 74.6 | 81.3 | 74 | 70 | 72.9349 | 82.4 | 71 | 76 | 78 | 81.4
| Mean |  | 72.8333 | 80.3 | 81.1667 | 74 | 73.3333 | 74.8903 | 81.5667 | 73 | 76.8333 | 77.1 | 82.7333
| Median |  | 73 | 80.9 | 81.1 | 74 | 70 | 74.9552 | 81.3 | 72 | 76.8 | 76.45 | 81.9
| Std Dev. |  | 1.3292 | 3.5367 | 0.92664 | 0 | 5.164 | 1.5672 | 1.8228 | 2.4495 | 0.81158 | 2.3238 | 2.3321
| Rel.Std.Dev. |  | 1.825% | 4.404% | 1.142% | 0% | 7.042% | 2.093% | 2.235% | 3.355% | 1.056% | 3.014% | 2.819%
| PDM3 |  | -5.398% | 4.3% | 5.426% | -3.883% | -4.749% | -2.726% | 5.945% | -5.182% | -0.2027% | 0.1436% | 7.461%
| Z-Score (Absolute) |  | 0.77 | 0.67 | 0.83 | 0.55 | 0.67 | 0.37 | 0.91 | 0.74 | 0 | 0.05 | 1.14
| Analytical results for Li in OREAS 553b (Certified Value 65 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 12 | 17 | 18 | 19 | 20 | 25 | 26 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [9 empty] | 0
1
| 1 | 1 | 65 | 65 | 61 | 66 | 70 | 64.5329 | 70 | 61 | 54 | 85 | 60 | 75.2 | 69.3 | [9 empty] | 1
| 1 | 2 | 66 | 63 | 60 | 65 | 70 | 63.7226 | 70 | 60 | 54 | 59 | 60 | 73.9 | 65.6 | [9 empty] | #N/A
| 1 | 3 | 65 | 64 | 59 | 65 | 60 | 64.502 | 70 | 60 | 52 | 60 | 60 | 80.2 | 83.5 | [9 empty] | 16
| 1 | 4 | 68 | 64 | 59 | 65 | 70 | 66.4641 | 70 | 61 | 51 | 57 | 60 | 76.1 | 63.2 | [9 empty] | 64.8582
| 1 | 5 | 63 | 64 | 60 | 66 | 60 | 66.265 | 70 | 61 | 52 | 63 | 60 | 71.9 | 67 | [9 empty] | 116
| 1 | 6 | 67 | 64 | 60 | 65 | 70 | 64.4237 | 70 | 61 | 53 | 59 | 70 | 75.5 | 61.7
| Mean |  | 65.6667 | 64 | 59.8333 | 65.3333 | 66.6667 | 64.9851 | 70 | 60.6667 | 52.6667 | 63.8333 | 61.6667 | 75.4667 | 68.3833
| Median |  | 65.5 | 64 | 60 | 65 | 70 | 64.5174 | 70 | 61 | 52.5 | 59.5 | 60 | 75.35 | 66.3
| Std Dev. |  | 1.7512 | 0.63246 | 0.75277 | 0.5164 | 5.164 | 1.1111 | 0 | 0.5164 | 1.2111 | 10.553 | 4.0825 | 2.7587 | 7.881
| Rel.Std.Dev. |  | 2.667% | 0.9882% | 1.258% | 0.7904% | 7.746% | 1.71% | 0% | 0.8512% | 2.299% | 16.53% | 6.62% | 3.656% | 11.52%
| PDM3 |  | 1.247% | -1.323% | -7.747% | 0.7326% | 2.788% | 0.1956% | 7.928% | -6.463% | -18.8% | -1.58% | -4.921% | 16.36% | 5.435%
| Z-Score (Absolute) |  | 0.14 | 0.2 | 1.05 | 0.07 | 0.34 | 0 | 1.02 | 0.88 | 2.5 | 0.23 | 0.67 | 2.13 | 0.69
| Analytical results for Lu in OREAS 553b (Certified Value 0.56 ppm) | [23 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 06 | 15 | 17 | 22 | 25 | 26 | [16 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [16 empty] | 2
3
| 1 | 1 | 0.51 | 0.57471 | 0.66 | 0.5 | 0.54 | 0.6 | [16 empty] | 1
| 1 | 2 | 0.46 | 0.5504 | 0.6 | < 0.5 | 0.62 | 0.6 | [16 empty] | 8
| 1 | 3 | 0.47 | 0.55951 | 0.59 | < 0.5 | 0.52 | 0.6 | [16 empty] | 16
| 1 | 4 | 0.53 | 0.56985 | 0.61 | < 0.5 | 0.53 | 0.6 | [16 empty] | 0.55961
| 1 | 5 | 0.46 | 0.5446 | 0.61 | < 0.5 | 0.51 | 0.6 | [16 empty] | 117
| 1 | 6 | 0.5 | 0.54914 | 0.59 | < 0.5 | 0.53 | 0.6
| Mean |  | 0.48833 | 0.55803 | 0.61 | 0.5 | 0.54167 | 0.6
| Median |  | 0.485 | 0.55495 | 0.605 | 0.5 | 0.53 | 0.6
| Std Dev. |  | 0.029269 | 0.012147 | 0.026077 |  | 0.039707 | 0
| Rel.Std.Dev. |  | 5.994% | 2.177% | 4.275% |  | 7.331% | 0%
| PDM3 |  | -12.74% | -0.2812% | 9.005% | -10.65% | -3.206% | 7.218%
| Z-Score (Absolute) |  | 0.75 | 0.1 | 0.74 | 3.16 | 0.1 | 0.61
| Analytical results for Mg in OREAS 553b (Certified Value 3.57 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 3.67 | 3.679 | 3.58 | 3.59 | 2.9337 | 3.42 | 3.65 | 3.35 | 3.6348 | 3.546 | 3.552 | 3.498 | 3.389 | 3.582 | 3.5379 | 3.58 | 3.67 | 3.4418 | 3.56 | 2.42 | 3.62 | 3.5 | 1
| 1 | 2 | 3.68 | 3.588 | 3.58 | 3.57 | 2.939 | 3.11 | 3.65 | 3.56 | 3.5395 | 3.564 | 3.6 | 3.437 | 3.371 | 3.649 | 3.5138 | 3.55 | 3.75 | 3.5223 | 3.57 | 2.44 | 3.66 | 3.54 | #N/A
| 1 | 3 | 3.62 | 3.528 | 3.58 | 3.49 | 2.935 | 3.24 | 3.65 | 3.44 | 3.6845 | 3.618 | 3.54 | 3.462 | 3.341 | 3.661 | 3.553 | 3.55 | 3.66 | 3.7691 | 3.58 | 4.43 | 3.74 | 3.54 | 16
| 1 | 4 | 3.62 | 3.63 | 3.56 | 3.49 | 2.9418 | 3 | 3.65 | 3.55 | 3.57 | 3.588 | 3.6 | 3.45 | 3.371 | 3.582 | 3.5171 | 3.55 | 3.56 | 3.5107 | 3.53 | 2.41 | 3.64 | 3.52 | 3.5675
| 1 | 5 | 3.69 | 3.588 | 3.6 | 3.53 | 2.9363 | 2.9 | 3.66 | 3.37 | 3.5988 | 3.576 | 3.522 | 3.6 | 3.389 | 3.612 | 3.5624 | 3.55 | 3.62 | 3.6421 | 3.64 | 4.44 | 3.62 | 3.53 | 118
| 1 | 6 | 3.66 | 3.582 | 3.53 | 3.43 | 2.9399 | 3.19 | 3.65 | 3.54 | 3.6432 | 3.655 | 3.528 | 3.6 | 3.377 | 3.661 | 3.5336 | 3.59 | 3.62 | 3.4732 | 3.65 | 2.42 | 3.64 | 3.41
| Mean |  | 3.6567 | 3.5992 | 3.5717 | 3.5167 | 2.9376 | 3.1433 | 3.6517 | 3.4683 | 3.6118 | 3.5912 | 3.557 | 3.5078 | 3.373 | 3.6245 | 3.5363 | 3.5617 | 3.6467 | 3.5599 | 3.5883 | 3.0933 | 3.6533 | 3.5067
| Median |  | 3.665 | 3.588 | 3.58 | 3.51 | 2.9377 | 3.15 | 3.65 | 3.49 | 3.6168 | 3.582 | 3.546 | 3.48 | 3.374 | 3.6305 | 3.5358 | 3.55 | 3.64 | 3.5165 | 3.575 | 2.43 | 3.64 | 3.525
| Std Dev. |  | 0.030111 | 0.050858 | 0.024014 | 0.058878 | 0.003116 | 0.18381 | 0.0040825 | 0.094534 | 0.052804 | 0.039529 | 0.034865 | 0.074228 | 0.017664 | 0.037506 | 0.019207 | 0.018348 | 0.06377 | 0.12314 | 0.047081 | 1.0393 | 0.045019 | 0.049666
| Rel.Std.Dev. |  | 0.8235% | 1.413% | 0.6723% | 1.674% | 0.1061% | 5.848% | 0.1118% | 2.726% | 1.462% | 1.101% | 0.9802% | 2.116% | 0.5237% | 1.035% | 0.5431% | 0.5152% | 1.749% | 3.459% | 1.312% | 33.6% | 1.232% | 1.416%
| PDM3 |  | 2.499% | 0.8875% | 0.1167% | -1.425% | -17.66% | -11.89% | 2.359% | -2.78% | 1.241% | 0.6633% | -0.2945% | -1.673% | -5.452% | 1.598% | -0.8747% | -0.1636% | 2.219% | -0.2141% | 0.5838% | -13.29% | 2.406% | -1.705%
| Z-Score (Absolute) |  | 1.21 | 0.49 | 0.14 | 0.56 | 7.85 | 5.26 | 1.15 | 1.16 | 0.64 | 0.38 | 0.05 | 0.67 | 2.36 | 0.8 | 0.31 | 0.01 | 1.08 | 0.01 | 0.35 | 5.89 | 1.17 | 0.68
| Analytical results for Mn in OREAS 553b (Certified Value 0.014 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | 3
3
| 1 | 1 | 0.0138 | 0.015 | 0.0142 | 0.014 | 0.013203 | 0.0165 | 0.017 | 0.013 | 0.012768 | 0.015 | 0.015 | 0.015 | 0.008 | 0.015 | 0.012935 | 0.0138 | 0.0137 | 0.0116 | 0.01 | 0.01214 | 0.003 | 0.0155 | 1
| 1 | 2 | 0.0141 | 0.015 | 0.0137 | 0.014 | 0.013572 | 0.015 | 0.0169 | 0.014 | 0.012421 | 0.015 | 0.015 | 0.015 | 0.008 | 0.015 | 0.012984 | 0.0135 | 0.0141 | 0.0117 | 0.01 | 0.01236 | 0.004 | 0.0139 | #N/A
| 1 | 3 | 0.0137 | 0.015 | 0.0148 | 0.014 | 0.013537 | 0.0155 | 0.017 | 0.012 | 0.012918 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.012728 | 0.0134 | 0.0136 | 0.0122 | 0.01 | 0.01225 | 0.004 | 0.0155 | 16
| 1 | 4 | 0.0137 | 0.015 | 0.0151 | 0.014 | 0.013293 | 0.0145 | 0.0169 | 0.014 | 0.012418 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.012692 | 0.0135 | 0.0138 | 0.0109 | 0.01 | 0.01219 | 0.005 | 0.0137 | 0.014047
| 1 | 5 | 0.0138 | 0.015 | 0.0158 | 0.013 | 0.013577 | 0.0139 | 0.0169 | 0.012 | 0.01327 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.013081 | 0.0136 | 0.0137 | 0.0121 | 0.01 | 0.01217 | 0.004 | 0.0155 | 119
| 1 | 6 | 0.0138 | 0.015 | 0.0133 | 0.013 | 0.013368 | 0.0159 | 0.017 | 0.013 | 0.012975 | 0.015 | 0.015 | 0.015 | 0.008 | 0.015 | 0.01276 | 0.0137 | 0.0136 | 0.0111 | 0.01 | 0.01224 | 0.005 | 0.014
| Mean |  | 0.013817 | 0.015 | 0.014483 | 0.013667 | 0.013425 | 0.015217 | 0.01695 | 0.013 | 0.012795 | 0.015 | 0.015 | 0.015 | 0.0115 | 0.015 | 0.012863 | 0.013583 | 0.01375 | 0.0116 | 0.01 | 0.012225 | 0.0041667 | 0.014683
| Median |  | 0.0138 | 0.015 | 0.0145 | 0.014 | 0.013453 | 0.01525 | 0.01695 | 0.013 | 0.012843 | 0.015 | 0.015 | 0.015 | 0.0115 | 0.015 | 0.012847 | 0.01355 | 0.0137 | 0.01165 | 0.01 | 0.012215 | 0.004 | 0.01475
| Std Dev. |  | 0.0001472 | 0 | 0.00092826 | 0.0005164 | 0.00015932 | 0.00094745 | 5.4772e-05 | 0.00089443 | 0.00033368 | 0 | 0 | 0 | 0.0038341 | 0 | 0.00015842 | 0.0001472 | 0.00018708 | 0.00052154 | 0 | 7.8166e-05 | 0.00075277 | 0.00089981
| Rel.Std.Dev. |  | 1.065% | 0% | 6.409% | 3.779% | 1.187% | 6.226% | 0.3231% | 6.88% | 2.608% | 0% | 0% | 0% | 33.34% | 0% | 1.232% | 1.084% | 1.361% | 4.496% | 0% | 0.6394% | 18.07% | 6.128%
| PDM3 |  | -1.642% | 6.782% | 3.104% | -2.71% | -4.429% | 8.325% | 20.66% | -7.455% | -8.915% | 6.782% | 6.782% | 6.782% | -18.13% | 6.782% | -8.429% | -3.303% | -2.116% | -17.42% | -28.81% | -12.97% | -70.34% | 4.528%
| Z-Score (Absolute) |  | 0.05 | 0.8 | 0.35 | 0.02 | 0.13 | 0.68 | 1.46 | 0.32 | 0.41 | 0.8 | 0.8 | 0.8 | 0.94 | 0.8 | 0.38 | 0.06 | 0.02 | 0.95 | 1.67 | 0.67 | 4.3 | 0.44
| Analytical results for Mo in OREAS 553b (Certified Value 16 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [10 empty] | 1
2
| 1 | 1 | 16 | 20 | 16 | 14 | < 10 | 18.3405 | 15.2262 | 15 | 15 | 8.2 | 16 | 15 | [10 empty] | 1
| 1 | 2 | 16 | 18 | 16 | 14 | < 10 | 17.5011 | 14.3435 | 15 | 16 | 8.2 | 19 | 15 | [10 empty] | 9
| 1 | 3 | 17 | 17 | 16 | 14 | < 10 | 18.1147 | 14.6289 | 15 | 17 | 8.3 | 18 | 16 | [10 empty] | 16
| 1 | 4 | 15 | 17 | 15 | 14 | < 10 | 17.6387 | 15.076 | 15 | 16 | 8.2 | 16 | 17 | [10 empty] | 15.9784
| 1 | 5 | 16 | 21 | 15 | 14 | < 10 | 18.1439 | 14.5112 | 16 | 17 | 8.2 | 15 | 17 | [10 empty] | 120
| 1 | 6 | 16 | 17 | 16 | 14 | < 10 | 17.9935 | 15.3877 | 15 | 17 | 8.2 | 12 | 16
| Mean |  | 16 | 18.3333 | 15.6667 | 14 |  | 17.9554 | 14.8622 | 15.1667 | 16.3333 | 8.2167 | 16 | 16
| Median |  | 16 | 17.5 | 16 | 14 |  | 18.0541 | 14.8524 | 15 | 16.5 | 8.2 | 16 | 16
| Std Dev. |  | 0.63246 | 1.7512 | 0.5164 | 0 |  | 0.32166 | 0.4245 | 0.40825 | 0.8165 | 0.040825 | 2.4495 | 0.89443
| Rel.Std.Dev. |  | 3.953% | 9.552% | 3.296% | 0% |  | 1.791% | 2.856% | 2.692% | 4.999% | 0.4969% | 15.31% | 5.59%
| PDM3 |  | 0.135% | 14.74% | -1.951% | -12.38% |  | 12.37% | -6.986% | -5.08% | 2.221% | -48.58% | 0.135% | 0.135%
| Z-Score (Absolute) |  | 0.14 | 2.06 | 0.14 | 1.51 | 8.92 | 1.75 | 0.8 | 0.55 | 0.41 | 6.27 | 0.14 | 0.14
| Analytical results for Na in OREAS 553b (Indicative Value 0.067 wt.%) | [23 empty] | Indicative
Na | Round | Replicate | Lab | Lab | [20 empty] | 1
| No | No | 02 | 22 | [20 empty] | wt.%
|  |  | BF*OES | PF*OES | [20 empty] | 3
3
| 1 | 1 | 0.052 | < 0.01 | [20 empty] | 1
| 1 | 2 | 0.082 | < 0.01 | [20 empty] | 20
| 1 | 3 | 0.059 | < 0.01 | [20 empty] | 16
| 1 | 4 | 0.089 | < 0.01 | [20 empty] | 0.066766
| 1 | 5 | 0.067 | < 0.01 | [20 empty] | 26
| 1 | 6 | 0.052 | < 0.01
| Mean |  | 0.066833
| Median |  | 0.063
| Std Dev. |  | 0.015639
| Rel.Std.Dev. |  | 23.4%
| PDM3 |  | 0.1015%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 553b (Certified Value 25.6 ppm) | [23 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 09 | 12 | 14 | 15 | 17 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 1
2
| 1 | 1 | 29.1 | 25.0056 | 29 | < 50 | 27.8012 | 25 | 24.3 | 22 | 28 | [13 empty] | 1
| 1 | 2 | 26.8 | 24.7488 | 33 | < 50 | 26.3436 | 24 | 24.9 | 23 | 26 | [13 empty] | 26
| 1 | 3 | 28 | 24.216 | 30 | < 50 | 25.7639 | 24 | 25.4 | 26 | 30 | [13 empty] | 16
| 1 | 4 | 27.2 | 24.9422 | 30 | < 50 | 25.2858 | 25 | 25.6 | 24 | 27 | [13 empty] | 25.5676
| 1 | 5 | 27.2 | 26.4118 | 28 | < 50 | 25.5744 | 25 | 25.4 | 29 | 31 | [13 empty] | 121
| 1 | 6 | 24.6 | 26.277 | 28 | < 50 | 25.2252 | 25 | 25.5 | 29 | 28
| Mean |  | 27.15 | 25.2669 | 29.6667 |  | 25.999 | 24.6667 | 25.1833 | 25.5 | 28.3333
| Median |  | 27.2 | 24.9739 | 29.5 |  | 25.6692 | 25 | 25.4 | 25 | 28
| Std Dev. |  | 1.4937 | 0.88063 | 1.8619 |  | 0.97051 | 0.5164 | 0.49565 | 3.0166 | 1.8619
| Rel.Std.Dev. |  | 5.501% | 3.485% | 6.276% |  | 3.733% | 2.094% | 1.968% | 11.83% | 6.571%
| PDM3 |  | 6.189% | -1.176% | 16.03% |  | 1.687% | -3.524% | -1.503% | -0.2643% | 10.82%
| Z-Score (Absolute) |  | 2.23 | 0.31 | 5.62 | 0.67 | 0.67 | 1.12 | 0.43 | 0 | 3.82
| Analytical results for Nd in OREAS 553b (Certified Value 65 ppm) | [23 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 0
1
| 1 | 1 | 70.7 | 63.3 | 64.282 | 67.6 | 57 | 62.66 | 64.6 | 68 | [14 empty] | 1
| 1 | 2 | 70 | 61.9 | 64.5212 | 63.7 | 58 | 62.89 | 65.6 | 63.1 | [14 empty] | 12
| 1 | 3 | 64.7 | 63.2 | 64.9197 | 65.2 | 62 | 62.38 | 64.8 | 65.5 | [14 empty] | 16
| 1 | 4 | 61.8 | 63.6 | 62.3507 | 64.7 | 58 | 63.89 | 63.4 | 64.9 | [14 empty] | 64.6418
| 1 | 5 | 68.8 | 62.8 | 62.7742 | 65.9 | 60 | 62.75 | 63.3 | 65.9 | [14 empty] | 122
| 1 | 6 | 55.2 | 63.1 | 64.4582 | 66.9 | 57 | 63.18 | 65.9 | 63.8
| Mean |  | 65.2 | 62.9833 | 63.8843 | 65.6667 | 58.6667 | 62.9583 | 64.6 | 65.2
| Median |  | 66.75 | 63.15 | 64.3701 | 65.55 | 58 | 62.82 | 64.7 | 65.2
| Std Dev. |  | 5.9676 | 0.59133 | 1.0535 | 1.4376 | 1.9664 | 0.52701 | 1.0826 | 1.7251
| Rel.Std.Dev. |  | 9.153% | 0.9389% | 1.649% | 2.189% | 3.352% | 0.8371% | 1.676% | 2.646%
| PDM3 |  | 0.8635% | -2.566% | -1.172% | 1.585% | -9.243% | -2.604% | -0.06468% | 0.8635%
| Z-Score (Absolute) |  | 0.58 | 0.77 | 0.22 | 0.87 | 3.39 | 0.78 | 0.22 | 0.58
| Analytical results for Ni in OREAS 553b (Certified Value 126 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 22 | 25 | 26 |  |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 0
0
| 1 | 1 | 104 | 130 | 140 | 80 | 134 | 110 | 126.4982 | 130 | 130 | 120 | 180 | 150 | 123 | 126 | 110 | 128.64 | 132 | 130 |  |  |  |  | 1
| 1 | 2 | 101 | 160 | 130 | 80 | 134 | 120 | 123.4285 | 130 | 130 | 120 | 210 | 140 | 124 | 126 | 110 | 129.13 | 125 | 132 |  |  |  |  | 29
| 1 | 3 | 108 | 120 | 130 | 80 | 134 | 110 | 127.2676 | 130 | 130 | 120 | 220 | 150 | 123 | 121 | 100 | 128.91 | 140 | 142 |  |  |  |  | 16
| 1 | 4 | 98 | 130 | 130 | 80 | 134 | 120 | 125.001 | 130 | 140 | 120 | 200 | 140 | 123 | 118 | 100 | 128.64 | 128 | 131 |  |  |  |  | 125.636
| 1 | 5 | 109 | 120 | 140 | 80 | 134 | 110 | 130.1566 | 110 | 130 | 190 | 240 | 140 | 122 | 120 | 110 | 129.04 | 143 | 139 |  |  |  |  | 123
| 1 | 6 | 99 | 120 | 130 | 80 | 134 | 120 | 129.2343 | 120 | 130 | 130 | 170 | 180 | 123 | 134 | 100 | 129.11 | 139 | 139
| Mean |  | 103.1667 | 130 | 133.3333 | 80 | 134 | 115 | 126.931 | 125 | 131.6667 | 133.3333 | 203.3333 | 150 | 123 | 124.1667 | 105 | 128.9117 | 134.5 | 135.5
| Median |  | 102.5 | 125 | 130 | 80 | 134 | 115 | 126.8829 | 130 | 130 | 120 | 205 | 145 | 123 | 123.5 | 105 | 128.975 | 135.5 | 135.5
| Std Dev. |  | 4.6224 | 15.4919 | 5.164 | 0 | 0 | 5.4772 | 2.5299 | 8.3666 | 4.0825 | 28.0476 | 25.8199 | 15.4919 | 0.63246 | 5.8109 | 5.4772 | 0.22409 | 7.2319 | 5.0892
| Rel.Std.Dev. |  | 4.481% | 11.92% | 3.873% | 0% | 0% | 4.763% | 1.993% | 6.693% | 3.101% | 21.04% | 12.7% | 10.33% | 0.5142% | 4.68% | 5.216% | 0.1738% | 5.377% | 3.756%
| PDM3 |  | -17.88% | 3.474% | 6.127% | -36.32% | 6.657% | -8.466% | 1.031% | -0.5062% | 4.8% | 6.127% | 61.84% | 19.39% | -2.098% | -1.17% | -16.43% | 2.607% | 7.055% | 7.851%
| Z-Score (Absolute) |  | 3.43 | 0.07 | 0.51 | 6.45 | 0.59 | 1.89 | 0.33 | 0.58 | 0.29 | 0.51 | 9.64 | 2.68 | 0.84 | 0.69 | 3.19 | 0.07 | 0.66 | 0.79
| Analytical results for P in OREAS 553b (Certified Value 0.037 wt.%) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 06 | 07 | 08 | 17 | 19 | 22 | 25 | 26 | [12 empty] | wt.%
|  |  | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.035 | 0.04 | 0.1 | 0.0927 | 0.03 | 0.04 | 0.0364 | 0.03 | 0.06 | 0.05 | [12 empty] | 1
| 1 | 2 | 0.035 | 0.04 | 0.1 | 0.0923 | 0.03 | 0.04 | 0.0369 | 0.04 | 0.06 | 0.03 | [12 empty] | #N/A
| 1 | 3 | 0.035 | 0.04 | 0.1 | 0.0923 | 0.03 | 0.04 | 0.0417 | 0.03 | 0.05 | 0.02 | [12 empty] | 16
| 1 | 4 | 0.035 | 0.04 | 0.1 | 0.0924 | 0.03 | 0.04 | 0.0368 | 0.03 | 0.03 | 0.04 | [12 empty] | 0.037287
| 1 | 5 | 0.035 | 0.04 | 0.08 | 0.0927 | 0.03 | 0.04 | 0.0398 | 0.03 | 0.04 | 0.04 | [12 empty] | 124
| 1 | 6 | 0.035 | 0.04 | 0.1 | 0.0926 | 0.03 | 0.03 | 0.0387 | 0.03 | 0.05 | 0.04
| Mean |  | 0.035 | 0.04 | 0.096667 | 0.0925 | 0.03 | 0.038333 | 0.038383 | 0.031667 | 0.048333 | 0.036667
| Median |  | 0.035 | 0.04 | 0.1 | 0.0925 | 0.03 | 0.04 | 0.0378 | 0.03 | 0.05 | 0.04
| Std Dev. |  | 0 | 0 | 0.008165 | 0.00018974 | 0 | 0.0040825 | 0.0020856 | 0.0040825 | 0.01169 | 0.010328
| Rel.Std.Dev. |  | 0% | 0% | 8.447% | 0.2051% | 0% | 10.65% | 5.434% | 12.89% | 24.19% | 28.17%
| PDM3 |  | -6.133% | 7.277% | 159.3% | 148.1% | -19.54% | 2.807% | 2.941% | -15.07% | 29.63% | -1.663%
| Z-Score (Absolute) |  | 0.46 | 0.22 | 7.76 | 7.2 | 1.11 | 0 | 0 | 0.89 | 1.33 | 0.22
| Analytical results for Pb in OREAS 553b (Certified Value 32.8 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 22 | 25 | 26 | [6 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | [6 empty] | 1
1
| 1 | 1 | < 100 | 13 | < 20 | 60 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 13 | 57 | < 300 | 17.7 | 11 | 14 | [6 empty] | 1
| 1 | 2 | < 100 | 12.8 | < 20 | 50 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 12 | 39 | < 300 | 17.7 | 11 | 13 | [6 empty] | 13
| 1 | 3 | < 100 | 10 | < 20 | 60 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 12 |  | < 300 | 17.7 | 12 | 14 | [6 empty] | 16
| 1 | 4 | < 100 | 13.5 | < 20 | 60 | < 20 | < 100 | < 100 | < 100 | < 100 | < 100 | 12 |  | < 300 | 17.7 | 11 | 13 | [6 empty] | 32.8276
| 1 | 5 | < 100 | 14.2 | < 20 | 50 | < 20 | < 100 | < 100 | < 100 | < 100 | 100 | 13 |  | < 300 | 17.7 | 11 | 13 | [6 empty] | 125
| 1 | 6 | < 100 | 13.4 | < 20 | 60 | < 20 | < 100 | < 100 | < 100 | < 100 | 100 | 13 |  | < 300 | 17.7 | 9 | 13
| Mean |  |  | 12.8167 |  | 56.6667 | [5 empty] | 100 | 12.5 | 48 |  | 17.7 | 10.8333 | 13.3333
| Median |  |  | 13.2 |  | 60 | [5 empty] | 100 | 12.5 | 48 |  | 17.7 | 11 | 13
| Std Dev. |  |  | 1.4621 |  | 5.164 | [5 empty] | 0 | 0.54772 | 12.7279 |  | 0 | 0.98319 | 0.5164
| Rel.Std.Dev. |  |  | 11.41% |  | 9.113% | [5 empty] | 0% | 4.382% | 26.52% |  | 0% | 9.076% | 3.873%
| PDM3 |  |  | -60.96% |  | 72.62% | [5 empty] | 204.6% | -61.92% | 46.22% |  | -46.08% | -67% | -59.38%
| Z-Score (Absolute) |  | 0.03 | 1 | 1.07 | 0.21 | 1.07 | 0.03 | 0.03 | 0.03 | 0.03 | 0.49 | 1.01 | 0.03 | 2.78 | 0.86 | 1.05 | 0.98
| Analytical results for Pr in OREAS 553b (Certified Value 18.1 ppm) | [23 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 06 | 15 | 17 | 22 | 25 | 26 | [15 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [15 empty] | 1
2
| 1 | 1 | 18.5 | 18.8 | 18.1047 | 19.14 | 17.7 | 17.7 | 18.61 | [15 empty] | 1
| 1 | 2 | 18.4 | 18.74 | 18.7871 | 17.99 | 17.7 | 18.4 | 17.75 | [15 empty] | 14
| 1 | 3 | 16.9 | 18.66 | 18.3387 | 18.33 | 17.4 | 17.3 | 17.95 | [15 empty] | 16
| 1 | 4 | 15 | 18.55 | 18.483 | 18.1 | 17.9 | 17.5 | 17.93 | [15 empty] | 18.131
| 1 | 5 | 18 | 18.81 | 18.2227 | 18.3 | 17.7 | 17.1 | 18.59 | [15 empty] | 126
| 1 | 6 | 15.2 | 18.4 | 18.2158 | 18.75 | 17.9 | 18 | 17.95
| Mean |  | 17 | 18.66 | 18.3586 | 18.435 | 17.7167 | 17.6667 | 18.13
| Median |  | 17.45 | 18.7 | 18.2807 | 18.315 | 17.7 | 17.6 | 17.95
| Std Dev. |  | 1.5786 | 0.16012 | 0.24609 | 0.43261 | 0.18348 | 0.4761 | 0.37181
| Rel.Std.Dev. |  | 9.286% | 0.8581% | 1.34% | 2.347% | 1.036% | 2.695% | 2.051%
| PDM3 |  | -6.238% | 2.918% | 1.256% | 1.677% | -2.285% | -2.561% | -0.005501%
| Z-Score (Absolute) |  | 1.84 | 0.86 | 0.37 | 0.5 | 0.67 | 0.76 | 0
| Analytical results for Rb in OREAS 553b (Certified Value 96 ppm) | [23 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 14 | 15 | 17 | 18 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 0
1
| 1 | 1 | 94.7 | 98.3 | 108 | 94 | 92.9891 | 95 | 100 | 96.2958 | 91.6 | 92.1 | 97 | 91.4 | [10 empty] | 1
| 1 | 2 | 94.1 | 98.5 | 107 | 94 | 90.4764 | 97.3 | 100 | 94.203 | 91 | 96.7 | 98 | 89.5 | [10 empty] | 15
| 1 | 3 | 94.8 | 99 | 106 | 94 | 92.093 | 96.7 | 100 | 93.5452 | 93.2 | 96.2 | 102 | 95.8 | [10 empty] | 16
| 1 | 4 | 95.7 | 99.4 | 102 | 93 | 90.8158 | 95.2 | 100 | 93.9085 | 94.3 | 93.7 | 97 | 92 | [10 empty] | 96.2434
| 1 | 5 | 93.9 | 102 | 101 | 93 | 94.2607 | 104.5 | 110 | 94.2321 | 93.1 | 93.2 | 96 | 94.4 | [10 empty] | 127
| 1 | 6 | 97.3 | 97.3 | 98.3 | 94 | 94.7144 | 104 | 100 | 95.9886 | 94.4 | 94.2 | 79 | 89.2
| Mean |  | 95.0833 | 99.0833 | 103.7167 | 93.6667 | 92.5582 | 98.7833 | 101.6667 | 94.6955 | 92.9333 | 94.35 | 94.8333 | 92.05
| Median |  | 94.75 | 98.75 | 104 | 94 | 92.5411 | 97 | 100 | 94.2175 | 93.15 | 93.95 | 97 | 91.7
| Std Dev. |  | 1.2561 | 1.5968 | 3.8473 | 0.5164 | 1.7504 | 4.3264 | 4.0825 | 1.1518 | 1.388 | 1.7762 | 8.0353 | 2.6319
| Rel.Std.Dev. |  | 1.321% | 1.612% | 3.709% | 0.5513% | 1.891% | 4.38% | 4.016% | 1.216% | 1.494% | 1.883% | 8.473% | 2.859%
| PDM3 |  | -1.205% | 2.951% | 7.765% | -2.677% | -3.829% | 2.639% | 5.635% | -1.608% | -3.439% | -1.967% | -1.465% | -4.357%
| Z-Score (Absolute) |  | 0.11 | 1.44 | 2.99 | 0.37 | 0.74 | 1.34 | 2.31 | 0.02 | 0.61 | 0.14 | 0.02 | 0.91
| Analytical results for Re in OREAS 553b (Certified Value < 0.1 ppm) | [23 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 04 | 06 | 09 | 25 | 26 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [17 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 1
| 1 | 2 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 16
| 1 | 3 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 16
| 1 | 4 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | [17 empty] | 128
| 1 | 6 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 553b (Certified Value 3.29 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 19 | 20 | 22 | 25 | 26 |  |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 3.47 | 3.27 | 3.37 | 3.206 | 3.5 | 3.42 | 2.75 | 3.5727 | 3.41 | 3.33 | 3.35 | 3.06 | 3.54 | 3.37 | 2.9692 | 3.18 | 3.48 | 2.69 | 3.35 |  |  |  | 1
| 1 | 2 | 3.47 | 3.22 | 3.41 | 3.2434 | 3.05 | 3.4 | 2.89 | 3.4909 | 3.39 | 3.35 | 3.32 | 3.06 | 3.59 | 3.43 | 3.1236 | 3.21 | 3.46 | 2.78 | 3.26 |  |  |  | 31
| 1 | 3 | 3.45 | 3.25 | 3.37 | 3.2308 | 3.17 | 3.41 | 2.77 | 3.5988 | 3.46 | 3.29 | 3.33 | 3.06 | 3.55 | 3.33 | 3.2512 | 3.14 | 3.49 | 2.82 | 3.26 |  |  |  | 16
| 1 | 4 | 3.39 | 3.2 | 3.31 | 3.2112 | 3.03 | 3.42 | 2.96 | 3.372 | 3.42 | 3.38 | 3.31 | 3.05 | 3.48 | 3.3 | 3.0598 | 3.18 | 3.47 | 2.86 | 3.29 |  |  |  | 3.2919
| 1 | 5 | 3.4 | 3.23 | 3.38 | 3.2823 | 2.94 | 3.41 | 2.76 | 3.5389 | 3.4 | 3.16 | 3.4 | 3.06 | 3.52 | 3.33 | 3.1978 | 3.26 | 3.45 | 2.78 | 3.27 |  |  |  | 129
| 1 | 6 | 3.46 | 3.32 | 3.42 | 3.2969 | 3.16 | 3.41 | 2.85 | 3.3528 | 3.47 | 3.27 | 3.43 | 3.09 | 3.52 | 3.3 | 2.8901 | 3.32 | 3.48 | 2.86 | 3.09
| Mean |  | 3.44 | 3.2483 | 3.3767 | 3.2451 | 3.1417 | 3.4117 | 2.83 | 3.4877 | 3.425 | 3.2967 | 3.3567 | 3.0633 | 3.5333 | 3.3433 | 3.082 | 3.215 | 3.4717 | 2.7983 | 3.2533
| Median |  | 3.455 | 3.24 | 3.375 | 3.2371 | 3.105 | 3.41 | 2.81 | 3.5149 | 3.415 | 3.31 | 3.34 | 3.06 | 3.53 | 3.33 | 3.0917 | 3.195 | 3.475 | 2.8 | 3.265
| Std Dev. |  | 0.035777 | 0.042622 | 0.038816 | 0.037289 | 0.19549 | 0.0075277 | 0.084617 | 0.1037 | 0.032711 | 0.077889 | 0.048028 | 0.013663 | 0.036697 | 0.049666 | 0.13698 | 0.065038 | 0.01472 | 0.064005 | 0.086872
| Rel.Std.Dev. |  | 1.04% | 1.312% | 1.15% | 1.149% | 6.223% | 0.2206% | 2.99% | 2.973% | 0.9551% | 2.363% | 1.431% | 0.446% | 1.039% | 1.486% | 4.444% | 2.023% | 0.424% | 2.287% | 2.67%
| PDM3 |  | 4.499% | -1.323% | 2.575% | -1.421% | -4.564% | 3.638% | -14.03% | 5.948% | 4.043% | 0.145% | 1.968% | -6.943% | 7.334% | 1.563% | -6.378% | -2.336% | 5.461% | -14.99% | -1.171%
| Z-Score (Absolute) |  | 0.75 | 0.25 | 0.42 | 0.27 | 0.81 | 0.6 | 2.45 | 1 | 0.67 | 0 | 0.32 | 1.23 | 1.24 | 0.25 | 1.13 | 0.43 | 0.92 | 2.62 | 0.23
| Analytical results for Sb in OREAS 553b (Certified Value 3.35 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 03 | 08 | 09 | 17 | 18 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 3.7 | 4 | < 50 | 3.1398 | 3.4 | 2.9 | < 5 | 3.3 | 3.9 | [13 empty] | 1
| 1 | 2 | 3.4 | 5 | < 50 | 2.8514 | 3.2 | 3.2 | < 5 | 3.5 | 3.4 | [13 empty] | 17
| 1 | 3 | 3.4 | 5 | < 50 | 3.2845 | 3.3 | 3.1 | < 5 | 3.6 | 3.9 | [13 empty] | 16
| 1 | 4 | 3.5 | 5 | < 50 | 3.1247 | 3.3 | 3.3 | < 5 | 3.4 | 3.9 | [13 empty] | 3.3495
| 1 | 5 | 3.4 | 5 | < 50 | 3.1124 | 3.3 | 3.2 | < 5 | 3.4 | 3.8 | [13 empty] | 130
| 1 | 6 | 3.2 | 4 | < 50 | 2.8308 | 3.4 | 3 | < 5 | 2.7 | 3.5
| Mean |  | 3.4333 | 4.6667 |  | 3.0573 | 3.3167 | 3.1167 |  | 3.3167 | 3.7333
| Median |  | 3.4 | 5 |  | 3.1185 | 3.3 | 3.15 |  | 3.4 | 3.85
| Std Dev. |  | 0.1633 | 0.5164 |  | 0.17872 | 0.075277 | 0.1472 |  | 0.31885 | 0.22509
| Rel.Std.Dev. |  | 4.756% | 11.07% |  | 5.846% | 2.27% | 4.723% |  | 9.614% | 6.029%
| PDM3 |  | 2.502% | 39.32% |  | -8.726% | -0.9815% | -6.952% |  | -0.9815% | 11.46%
| Z-Score (Absolute) |  | 0.3 | 3.51 | 56.36 | 0.67 | 0 | 0.52 | 2.12 | 0 | 1.08
| Analytical results for Sc in OREAS 553b (Certified Value 12.2 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 04 | 07 | 08 | 09 | 17 | 18 | 19 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | [10 empty] | 1
2
| 1 | 1 | 11 | 14 | < 20 | 12 | 13 | < 20 | 13 | 11 | 11 | 12.7 | 9 | 13 | [10 empty] | 1
| 1 | 2 | 10 | 14 | < 20 | 12 | 12 | < 20 | 13 | 11 | 11 | 12.4 | 8 | 13 | [10 empty] | 18
| 1 | 3 |  | 14 | < 20 | 12 | 12 | < 20 | 13 | 12 | 12 | 12.6 | 10 | 13 | [10 empty] | 16
| 1 | 4 | 11 | 14 | < 20 | 12 | 13 | < 20 | 12 | 11 | 11 | 12.6 | 7 | 13 | [10 empty] | 12.2481
| 1 | 5 | 11 | 14 | < 20 | 12 | 13 | < 20 | 12 | 12 | 12 | 12.7 | 10 | 13 | [10 empty] | 131
| 1 | 6 | 11 | 14 | < 20 | 12 | 12 | < 20 | 13 | 11 | 11 | 12.6 | 8 | 13
| Mean |  | 10.8 | 14 |  | 12 | 12.5 |  | 12.6667 | 11.3333 | 11.3333 | 12.6 | 8.6667 | 13
| Median |  | 11 | 14 |  | 12 | 12.5 |  | 13 | 11 | 11 | 12.6 | 8.5 | 13
| Std Dev. |  | 0.44721 | 0 |  | 0 | 0.54772 |  | 0.5164 | 0.5164 | 0.5164 | 0.10954 | 1.2111 | 0
| Rel.Std.Dev. |  | 4.141% | 0% |  | 0% | 4.382% |  | 4.077% | 4.556% | 4.556% | 0.8694% | 13.97% | 0%
| PDM3 |  | -11.82% | 14.3% |  | -2.026% | 2.056% |  | 3.417% | -7.469% | -7.469% | 2.873% | -29.24% | 6.138%
| Z-Score (Absolute) |  | 0.6 | 1.63 | 1.16 | 0.23 | 0.58 | 1.16 | 0.7 | 0.23 | 0.23 | 0.65 | 2.09 | 0.93
| Analytical results for Se in OREAS 553b (Indicative Value 9.75 ppm) | [23 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 03 | 04 | 09 | 25 | 26 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [17 empty] | 2
2
| 1 | 1 | 16 | < 20 | < 20 | 7 | < 50 | [17 empty] | 1
| 1 | 2 | 13 | < 20 | < 20 | 9 | < 50 | [17 empty] | 21
| 1 | 3 | 12 | < 20 | < 20 | 6 | < 50 | [17 empty] | 16
| 1 | 4 | 10 | < 20 | < 20 | 9 | < 50 | [17 empty] | 9.75
| 1 | 5 | 13 | < 20 | < 20 | 5 | < 50 | [17 empty] | 27
| 1 | 6 | 11 | < 20 | < 20 | 6 | < 50
| Mean |  | 12.5 |  |  | 7
| Median |  | 12.5 |  |  | 6.5
| Std Dev. |  | 2.0736 |  |  | 1.6733
| Rel.Std.Dev. |  | 16.59% |  |  | 23.9%
| PDM3 |  | 28.21% |  |  | -28.21%
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 0.81 | 4.05
| Analytical results for Si in OREAS 553b (Certified Value 25.17 wt.%) | [23 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 25 | 26 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 2
3
| 1 | 1 | 25.9 | 24.1 | 25.6 | 25.1 | 23.8412 | 27.8 | 25.28 | 25.87 | 25.0565 | 24.68 | 25.01 | 25.38 | 25.38 | 25.9 | 24.6 | 17.71 | 26.7 | [5 empty] | 1
| 1 | 2 | 25.9 | 22.9 | 25.7 | 24.9 | 23.9046 | 24.6 | 25.27 | 26.26 | 24.3624 | 24.59 | 25.29 | 25.2 | 25.2 | 26.08 | 24.3 | 20.39 | 26.9 | [5 empty] | #N/A
| 1 | 3 | 25.5 | 25 | 25.4 | 24.3 | 23.7627 | 25.2 | 25.24 | 26.53 | 24.3579 | 25.1 | 24.92 | 25.29 | 25.29 | 26.08 | 24.3 | 18.92 | 27 | [5 empty] | 16
| 1 | 4 | 25.4 | 25 | 25.4 | 24.4 | 23.876 | 24 | 25.29 | 26.18 | 23.7615 | 24.82 | 25.29 | 25.2 | 25.38 | 25.38 | 24.4 | 19.45 | 26.6 | [5 empty] | 25.1713
| 1 | 5 | 25.8 | 24.6 | 25.3 | 24.6 | 23.8043 | 23.5 | 25.27 | 25.91 | 26.0451 | 24.92 | 24.82 | 25.48 | 25.29 | 25.8 | 24.3 | 18.71 | 26.8 | [5 empty] | 132
| 1 | 6 | 25.7 | 24.7 | 24.5 | 24 | 23.8381 | 26 | 25.21 | 26.44 | 26.3056 | 25.2 | 25.01 | 25.48 | 25.2 | 26.08 | 24.8 | 20.29 | 25.9
| Mean |  | 25.7 | 24.3833 | 25.3167 | 24.55 | 23.8378 | 25.1833 | 25.26 | 26.1983 | 24.9815 | 24.885 | 25.0567 | 25.3383 | 25.29 | 25.8867 | 24.45 | 19.245 | 26.65
| Median |  | 25.75 | 24.65 | 25.4 | 24.5 | 23.8396 | 24.9 | 25.27 | 26.22 | 24.7094 | 24.87 | 25.01 | 25.335 | 25.29 | 25.99 | 24.35 | 19.185 | 26.75
| Std Dev. |  | 0.20976 | 0.79854 | 0.42622 | 0.40373 | 0.050381 | 1.5549 | 0.029665 | 0.26962 | 1.015 | 0.23662 | 0.19387 | 0.12844 | 0.080498 | 0.27442 | 0.20736 | 1.0191 | 0.3937
| Rel.Std.Dev. |  | 0.8162% | 3.275% | 1.684% | 1.645% | 0.2113% | 6.174% | 0.1174% | 1.029% | 4.063% | 0.9509% | 0.7737% | 0.5069% | 0.3183% | 1.06% | 0.8481% | 5.296% | 1.477%
| PDM3 |  | 2.1% | -3.131% | 0.5774% | -2.468% | -5.298% | 0.04769% | 0.3523% | 4.08% | -0.7541% | -1.138% | -0.4555% | 0.6635% | 0.4714% | 2.842% | -2.866% | -23.54% | 5.874%
| Z-Score (Absolute) |  | 0.67 | 1.04 | 0.17 | 0.83 | 1.76 | 0 | 0.1 | 1.32 | 0.26 | 0.39 | 0.17 | 0.2 | 0.14 | 0.92 | 0.96 | 7.75 | 1.91
| Analytical results for Sm in OREAS 553b (Certified Value 8.9 ppm) | [23 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 8.5 | 9.1 | 8 | 9.2685 | 9.8 | 8.3 | 9.2 | 9.9 | [14 empty] | 1
| 1 | 2 | 8.6 | 9.1 | 8 | 8.9448 | 9.7 | 8.8 | 10 | 9.1 | [14 empty] | 5
| 1 | 3 | 8 | 8.9 | 8 | 9.0372 | 9.5 | 8.6 | 9.1 | 9.3 | [14 empty] | 16
| 1 | 4 | 7.6 | 8.8 | 8 | 9.1706 | 9.6 | 8.7 | 9.5 | 9.4 | [14 empty] | 8.9016
| 1 | 5 | 9.2 | 8.8 | 7.9 | 8.5566 | 9.6 | 8.7 | 9 | 10 | [14 empty] | 133
| 1 | 6 | 7.1 | 8.8 | 8 | 9.1975 | 10 | 8.8 | 9 | 9.1
| Mean |  | 8.1667 | 8.9167 | 7.9833 | 9.0292 | 9.7 | 8.65 | 9.3 | 9.4667
| Median |  | 8.25 | 8.85 | 8 | 9.1039 | 9.65 | 8.7 | 9.15 | 9.35
| Std Dev. |  | 0.75542 | 0.1472 | 0.040825 | 0.2593 | 0.17889 | 0.18708 | 0.38987 | 0.39328
| Rel.Std.Dev. |  | 9.25% | 1.651% | 0.5114% | 2.872% | 1.844% | 2.163% | 4.192% | 4.154%
| PDM3 |  | -8.256% | 0.1696% | -10.32% | 1.434% | 8.97% | -2.826% | 4.476% | 6.348%
| Z-Score (Absolute) |  | 1.32 | 0.09 | 1.63 | 0.09 | 1.19 | 0.53 | 0.54 | 0.81
| Analytical results for Sn in OREAS 553b (Indicative Value 2.86 ppm) | [23 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 03 | 08 | 09 | 12 | 14 | 15 | 17 | 20 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | [10 empty] | 2
2
| 1 | 1 | < 50 | 6.2 | < 50 | < 100 | < 5 | 6 | 2.8637 | 3 | < 50 | < 5 | 3 | < 10 | [10 empty] | 1
| 1 | 2 | < 50 | 3.8 | < 50 | < 100 | < 5 | < 5 | 2.8001 | 3 | < 50 | < 5 | 2.5 | < 10 | [10 empty] | 22
| 1 | 3 | < 50 | 4 | < 50 | < 100 | < 5 | < 5 | 2.5428 | 3 | < 50 | < 5 | 3.8 | < 10 | [10 empty] | 16
| 1 | 4 | < 50 | 4.1 | < 50 | < 100 | < 5 | < 5 | 2.5656 | 3 | < 50 | < 5 | 3.2 | < 10 | [10 empty] | 2.8557
| 1 | 5 | < 50 | 4 | < 50 | < 100 | < 5 | < 5 | 2.6655 | 3 | < 50 | < 5 | 3.9 | < 10 | [10 empty] | 28
| 1 | 6 | < 50 | 4.4 | < 50 | < 100 | < 5 | 5 | 2.6003 | 3 | < 50 | < 5 | 3.1 | < 10
| Mean |  |  | 4.4167 |  |  |  | 5.5 | 2.673 | 3 |  |  | 3.25
| Median |  |  | 4.05 |  |  |  | 5.5 | 2.6329 | 3 |  |  | 3.15
| Std Dev. |  |  | 0.89536 |  |  |  | 0.70711 | 0.13141 | 0 |  |  | 0.5244
| Rel.Std.Dev. |  |  | 20.27% |  |  |  | 12.86% | 4.916% | 0% |  |  | 16.14%
| PDM3 |  |  | 54.66% |  |  |  | 92.59% | -6.4% | 5.051% |  |  | 13.81%
| Z-Score (Absolute) |  | 10.34 | 0.23 | 10.34 | 22.63 | 0.72 | 0.23 | 0.63 | 0.47 | 10.34 | 0.72 | 0.35 | 0.51
| Analytical results for Sr in OREAS 553b (Certified Value 70 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [7 empty] | 0
1
| 1 | 1 | 72 | 70 | 83 | 70 | 95 | 43 | 70 | 73.5339 | 70.1373 | 71 | 71 | 65 | 20.1 | 68 | 71.3 | [7 empty] | 1
| 1 | 2 | 71 | 60 | 80 | 70 | 85 | 43 | 70 | 70.2838 | 72.0366 | 69 | 69 | 66 | 20.2 | 68 | 67.1 | [7 empty] | #N/A
| 1 | 3 | 72 | 80 | 80 | 70 | 90 | 44 | 60 | 73.4972 | 67.5315 | 70 | 69 | 71 | 20.1 | 73 | 70.1 | [7 empty] | 16
| 1 | 4 | 74 | 70 | 77 | 70 | 80 | 44 | 70 | 70.2087 | 74.9463 | 70 | 65 | 66 | 20.2 | 68 | 67.8 | [7 empty] | 69.6557
| 1 | 5 | 71 | 80 | 79 | 70 | 75 | 44 | 60 | 75.0836 | 66.2342 | 70 | 70 | 69 | 20.1 | 66 | 72.5 | [7 empty] | 134
| 1 | 6 | 77 | 70 | 70 | 70 | 89 | 44 | 70 | 73.5453 | 69.4404 | 71 | 69 | 66 | 20.2 | 64 | 66
| Mean |  | 72.8333 | 71.6667 | 78.1667 | 70 | 85.6667 | 43.6667 | 66.6667 | 72.6921 | 70.0544 | 70.1667 | 68.8333 | 67.1667 | 20.15 | 67.8333 | 69.1333
| Median |  | 72 | 70 | 79.5 | 70 | 87 | 44 | 70 | 73.5155 | 69.7888 | 70 | 69 | 66 | 20.15 | 68 | 68.95
| Std Dev. |  | 2.3166 | 7.5277 | 4.446 | 0 | 7.2572 | 0.5164 | 5.164 | 1.9885 | 3.137 | 0.75277 | 2.0412 | 2.3166 | 0.054772 | 2.9944 | 2.5571
| Rel.Std.Dev. |  | 3.181% | 10.5% | 5.688% | 0% | 8.471% | 1.183% | 7.746% | 2.736% | 4.478% | 1.073% | 2.965% | 3.449% | 0.2718% | 4.414% | 3.699%
| PDM3 |  | 4.562% | 2.887% | 12.22% | 0.4942% | 22.99% | -37.31% | -4.291% | 4.359% | 0.5723% | 0.7335% | -1.181% | -3.573% | -71.07% | -2.616% | -0.75%
| Z-Score (Absolute) |  | 0.71 | 0.42 | 2.05 | 0 | 3.92 | 6.6 | 0.83 | 0.67 | 0.01 | 0.04 | 0.29 | 0.71 | 12.49 | 0.54 | 0.22
| Analytical results for Ta in OREAS 553b (Certified Value 2.01 ppm) | [23 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 14 | 15 | 17 | 18 | 22 | 25 | 26 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [9 empty] | 2
2
| 1 | 1 | 1.7 | 2.4 | 2.8 | 1.7 | 2.3817 | 1.7 | < 10 | 2.0929 | 2.1 | 1.9 | 2.1 | 1.88 | 1.5 | [9 empty] | 1
| 1 | 2 | 2.2 | 2.2 | 2.7 | 1.6 | 2.1809 | 1.7 | < 10 | 2.1219 | 2 | 2.2 | 2.1 | 2.04 | 1.2 | [9 empty] | 7
| 1 | 3 | 1.9 | 2.4 | 2.8 | 1.7 | 2.121 | 2 | < 10 | 1.9434 | 1.9 | 1.8 | 2.1 | 1.91 | 2.5 | [9 empty] | 16
| 1 | 4 | 2.1 | 2.6 | 2.9 | 1.6 | 2.0976 | 1.8 | < 10 | 2.0888 | 2 | 1.9 | 2.1 | 1.97 | 1.2 | [9 empty] | 2.0098
| 1 | 5 | 2.4 | 2.5 | 2.9 | 1.7 | 2.3915 | 2.5 | < 10 | 2.1374 | 2 | 1.9 | 2.1 | 1.68 | 2.2 | [9 empty] | 135
| 1 | 6 | 2 | 2.2 | 2.8 | 1.7 | 2.4958 | 2.5 | < 10 | 1.8859 | 2.1 | 2 | 2.1 | 1.63 | 1.8
| Mean |  | 2.05 | 2.3833 | 2.8167 | 1.6667 | 2.2781 | 2.0333 |  | 2.0451 | 2.0167 | 1.95 | 2.1 | 1.8517 | 1.7333
| Median |  | 2.05 | 2.4 | 2.8 | 1.7 | 2.2813 | 1.9 |  | 2.0908 | 2 | 1.9 | 2.1 | 1.895 | 1.65
| Std Dev. |  | 0.2429 | 0.16021 | 0.075277 | 0.05164 | 0.16596 | 0.37771 |  | 0.10419 | 0.075277 | 0.13784 | 0 | 0.16266 | 0.53541
| Rel.Std.Dev. |  | 11.85% | 6.722% | 2.673% | 3.098% | 7.285% | 18.58% |  | 5.095% | 3.733% | 7.069% | 0% | 8.784% | 30.89%
| PDM3 |  | 1.999% | 18.58% | 40.14% | -17.07% | 13.35% | 1.169% |  | 1.753% | 0.3402% | -2.977% | 4.486% | -7.869% | -13.76%
| Z-Score (Absolute) |  | 0.02 | 1.18 | 2.69 | 1.32 | 0.81 | 0.04 | 10.3 | 0 | 0.1 | 0.33 | 0.19 | 0.67 | 1.09
| Analytical results for Tb in OREAS 553b (Certified Value 1.18 ppm) | [23 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 1.3 | 1.19 | 1.13 | 1.2884 | 1.31 | 1 | 1.12 | 1.28 | [14 empty] | 1
| 1 | 2 | 1.3 | 1.17 | 1.12 | 1.2315 | 1.23 | 1 | 1.16 | 1.28 | [14 empty] | 8
| 1 | 3 | 1.2 | 1.15 | 1.17 | 1.2582 | 1.2 | 1 | 1.1 | 1.37 | [14 empty] | 16
| 1 | 4 | 0.9 | 1.22 | 1.18 | 1.1435 | 1.23 | 1 | 1.19 | 1.31 | [14 empty] | 1.1832
| 1 | 5 | 1.2 | 1.19 | 1.12 | 1.1316 | 1.22 | 1 | 1.07 | 1.35 | [14 empty] | 136
| 1 | 6 | 1 | 1.16 | 1.17 | 1.1406 | 1.26 | 1 | 1.14 | 1.34
| Mean |  | 1.15 | 1.18 | 1.1483 | 1.199 | 1.2417 | 1 | 1.13 | 1.3217
| Median |  | 1.2 | 1.18 | 1.15 | 1.1875 | 1.23 | 1 | 1.13 | 1.325
| Std Dev. |  | 0.16432 | 0.025298 | 0.027869 | 0.06868 | 0.038687 | 0 | 0.042895 | 0.037639
| Rel.Std.Dev. |  | 14.29% | 2.144% | 2.427% | 5.728% | 3.116% | 0% | 3.796% | 2.848%
| PDM3 |  | -2.803% | -0.2672% | -2.944% | 1.336% | 4.945% | -15.48% | -4.493% | 11.71%
| Z-Score (Absolute) |  | 0.29 | 0.29 | 0.33 | 0.66 | 1.5 | 3.23 | 0.68 | 3.06
| Analytical results for Te in OREAS 553b (Indicative Value < 1 ppm) | [23 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 03 | 04 | 09 | 25 | 26 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [17 empty] | 2
2
| 1 | 1 | < 6 | < 1 | < 1 | 0.1 | < 1 | [17 empty] | 1
| 1 | 2 | < 6 | < 1 | < 1 | < 0.1 | < 1 | [17 empty] | 23
| 1 | 3 | < 6 | < 1 | < 1 | < 0.1 | < 1 | [17 empty] | 16
| 1 | 4 | < 6 | < 1 | < 1 | < 0.1 | < 1 | [17 empty] | < 1
| 1 | 5 | < 6 | < 1 | < 1 | 0.1 | < 1 | [17 empty] | 29
| 1 | 6 | < 6 | < 1 | < 1 | 0.1 | < 1
| Mean |  |  |  |  | 0.1
| Median |  |  |  |  | 0.1
| Std Dev. |  |  |  |  | 1.6997e-17
| Rel.Std.Dev. |  |  |  |  | 1.7e-14%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Th in OREAS 553b (Certified Value 16 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 12 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [10 empty] | 1
2
| 1 | 1 | 16 | 16.5 | 16.1 | 13.6 | 16.2419 | 14.4 | 16.3566 | 16.4 | 15.3 | 15.9 | 15.9 | 16 | [10 empty] | 1
| 1 | 2 | 15.8 | 15.6 | 16.1 | 13.2 | 15.7586 | 15.8 | 16.6134 | 16 | 16 | 15.7 | 15.5 | 15.2 | [10 empty] | 10
| 1 | 3 | 15.9 | 15.8 | 16.3 | 13.3 | 16.0369 | 14.7 | 15.7188 | 16.2 | 15.6 | 15.4 | 16.8 | 15.7 | [10 empty] | 16
| 1 | 4 | 16 | 15.2 | 16.1 | 13.8 | 15.8288 | 15.9 | 16.2944 | 15.9 | 15.4 | 15.7 | 16.2 | 15.7 | [10 empty] | 15.9533
| 1 | 5 | 16.3 | 16.3 | 16.1 | 13.4 | 16.3856 | 13.2 | 16.1964 | 16.4 | 15.4 | 15.5 | 15.3 | 16.5 | [10 empty] | 137
| 1 | 6 | 16.6 | 13.9 | 16.1 | 13.2 | 16.1989 | 13.6 | 15.8495 | 16.7 | 15.5 | 15.7 | 12.6 | 15.6
| Mean |  | 16.1 | 15.55 | 16.1333 | 13.4167 | 16.0751 | 14.6 | 16.1715 | 16.2667 | 15.5333 | 15.65 | 15.3833 | 15.7833
| Median |  | 16 | 15.7 | 16.1 | 13.35 | 16.1179 | 14.55 | 16.2454 | 16.3 | 15.45 | 15.7 | 15.7 | 15.7
| Std Dev. |  | 0.29665 | 0.93541 | 0.08165 | 0.24014 | 0.2457 | 1.1082 | 0.33282 | 0.29439 | 0.25033 | 0.17607 | 1.4634 | 0.43551
| Rel.Std.Dev. |  | 1.843% | 6.016% | 0.5061% | 1.79% | 1.528% | 7.59% | 2.058% | 1.81% | 1.612% | 1.125% | 9.513% | 2.759%
| PDM3 |  | 0.9194% | -2.528% | 1.128% | -15.9% | 0.7634% | -8.483% | 1.368% | 1.964% | -2.633% | -1.901% | -3.573% | -1.066%
| Z-Score (Absolute) |  | 0.7 | 0.3 | 0.76 | 4.18 | 0.65 | 2.03 | 0.83 | 1 | 0.33 | 0.12 | 0.61 | 0.12
| Analytical results for Ti in OREAS 553b (Certified Value 0.461 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.49 | 0.48 | 0.48 | 0.48 | 0.49 | 0.48 | 0.44 | 0.463 | 0.468 | 0.456 | 0.456 | 0.444 | 0.462 | 0.46842 | 0.46 | 0.48 | 0.4196 | 0.44 | 0.11 | 0.45 | 0.48 |  | 1
| 1 | 2 | 0.5 | 0.468 | 0.48 | 0.48 | 0.44 | 0.47 | 0.46 | 0.45111 | 0.492 | 0.456 | 0.456 | 0.438 | 0.474 | 0.46931 | 0.45 | 0.49 | 0.428 | 0.45 | 0.12 | 0.47 | 0.47 |  | 26
| 1 | 3 | 0.49 | 0.462 | 0.47 | 0.48 | 0.46 | 0.47 | 0.45 | 0.46531 | 0.474 | 0.456 | 0.45 | 0.438 | 0.468 | 0.4559 | 0.46 | 0.47 | 0.4589 | 0.45 | 0.11 | 0.45 | 0.48 |  | 16
| 1 | 4 | 0.48 | 0.468 | 0.47 | 0.47 | 0.43 | 0.47 | 0.46 | 0.45307 | 0.468 | 0.462 | 0.45 | 0.438 | 0.456 | 0.47054 | 0.46 | 0.46 | 0.4253 | 0.44 | 0.11 | 0.44 | 0.47 |  | 0.46085
| 1 | 5 | 0.49 | 0.462 | 0.47 | 0.47 | 0.41 | 0.48 | 0.44 | 0.45997 | 0.468 | 0.45 | 0.468 | 0.444 | 0.462 | 0.46567 | 0.46 | 0.47 | 0.4415 | 0.46 | 0.11 | 0.41 | 0.47 |  | 138
| 1 | 6 | 0.49 | 0.462 | 0.47 | 0.46 | 0.46 | 0.47 | 0.45 | 0.46823 | 0.474 | 0.45 | 0.468 | 0.438 | 0.468 | 0.46517 | 0.46 | 0.47 | 0.4235 | 0.46 | 0.11 | 0.43 | 0.45
| Mean |  | 0.49 | 0.467 | 0.47333 | 0.47333 | 0.44833 | 0.47333 | 0.45 | 0.46012 | 0.474 | 0.455 | 0.458 | 0.44 | 0.465 | 0.46584 | 0.45833 | 0.47333 | 0.4328 | 0.45 | 0.11167 | 0.44167 | 0.47
| Median |  | 0.49 | 0.465 | 0.47 | 0.475 | 0.45 | 0.47 | 0.45 | 0.46149 | 0.471 | 0.456 | 0.456 | 0.438 | 0.465 | 0.46705 | 0.46 | 0.47 | 0.42665 | 0.45 | 0.11 | 0.445 | 0.47
| Std Dev. |  | 0.0063246 | 0.0070143 | 0.005164 | 0.008165 | 0.027869 | 0.005164 | 0.0089443 | 0.0068107 | 0.0092952 | 0.0045166 | 0.0081976 | 0.0030984 | 0.0062929 | 0.0052942 | 0.0040825 | 0.010328 | 0.014812 | 0.0089443 | 0.0040825 | 0.020412 | 0.010954
| Rel.Std.Dev. |  | 1.291% | 1.502% | 1.091% | 1.725% | 6.216% | 1.091% | 1.988% | 1.48% | 1.961% | 0.9927% | 1.79% | 0.7042% | 1.353% | 1.137% | 0.8907% | 2.182% | 3.422% | 1.988% | 3.656% | 4.622% | 2.331%
| PDM3 |  | 6.325% | 1.334% | 2.708% | 2.708% | -2.717% | 2.708% | -2.355% | -0.1595% | 2.853% | -1.27% | -0.619% | -4.525% | 0.9% | 1.081% | -0.5466% | 2.708% | -6.087% | -2.355% | -75.77% | -4.163% | 1.985%
| Z-Score (Absolute) |  | 1.71 | 0.37 | 0.76 | 0.76 | 0.67 | 0.76 | 0.58 | 0 | 0.77 | 0.32 | 0.14 | 1.17 | 0.26 | 0.33 | 0.1 | 0.76 | 1.56 | 0.58 | 19.94 | 1.06 | 0.57
| Analytical results for Tl in OREAS 553b (Certified Value 0.73 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 17 | 18 | 22 | 25 | 26 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [12 empty] | 2
3
| 1 | 1 | 0.7 | 0.7 | 0.7 | 0.7 | 0.64705 | 0.8 | 0.7 | < 1 | 0.6 | 0.8 | [12 empty] | 1
| 1 | 2 | 0.8 | 0.7 | 0.7 | 0.7 | 0.66396 | 0.7 | 0.7 | < 1 | 0.7 | 0.8 | [12 empty] | 27
| 1 | 3 | 0.8 | 0.8 | 0.7 | 0.7 | 0.63155 | 0.8 | 0.8 | < 1 | 0.7 | 0.8 | [12 empty] | 16
| 1 | 4 | 0.7 | 0.8 | 0.7 | 0.7 | 0.6259 | 0.8 | 0.7 | < 1 | 0.7 | 0.7 | [12 empty] | 0.7292
| 1 | 5 | 0.8 | 0.7 | 0.7 | 0.7 | 0.67373 | 0.8 | 0.7 | < 1 | 0.7 | 0.9 | [12 empty] | 139
| 1 | 6 | 0.9 | 0.7 | 0.7 | 0.7 | 0.6548 | 0.8 | 0.8 | < 1 | 0.5 | 0.8
| Mean |  | 0.78333 | 0.73333 | 0.7 | 0.7 | 0.6495 | 0.78333 | 0.73333 |  | 0.65 | 0.8
| Median |  | 0.8 | 0.7 | 0.7 | 0.7 | 0.65092 | 0.8 | 0.7 |  | 0.7 | 0.8
| Std Dev. |  | 0.075277 | 0.05164 | 1.2162e-16 | 1.2162e-16 | 0.01849 | 0.040825 | 0.05164 |  | 0.083666 | 0.063246
| Rel.Std.Dev. |  | 9.61% | 7.042% | 1.737e-14% | 1.737e-14% | 2.847% | 5.212% | 7.042% |  | 12.87% | 7.906%
| PDM3 |  | 7.423% | 0.5664% | -4.005% | -4.005% | -10.93% | 7.423% | 0.5664% |  | -10.86% | 9.709%
| Z-Score (Absolute) |  | 0.67 | 0.17 | 0.17 | 0.17 | 0.68 | 0.67 | 0.17 | 2.19 | 0.67 | 0.84
| Analytical results for Tm in OREAS 553b (Certified Value 0.57 ppm) | [23 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 22 | 25 | 26 | [14 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | [14 empty] | 2
3
| 1 | 1 | 0.6 | 0.53 | 0.51 | 0.58311 | 0.64 | < 1 | 0.59 | 0.64 | [14 empty] | 1
| 1 | 2 | 0.6 | 0.53 | 0.49 | 0.59371 | 0.61 | < 1 | 0.67 | 0.6 | [14 empty] | 12
| 1 | 3 | 0.6 | 0.53 | 0.5 | 0.58442 | 0.59 | < 1 | 0.53 | 0.64 | [14 empty] | 16
| 1 | 4 | 0.5 | 0.56 | 0.52 | 0.59076 | 0.6 | < 1 | 0.59 | 0.6 | [14 empty] | 0.57021
| 1 | 5 | 0.6 | 0.52 | 0.47 | 0.55638 | 0.61 | < 1 | 0.53 | 0.69 | [14 empty] | 140
| 1 | 6 | 0.5 | 0.54 | 0.49 | 0.54052 | 0.61 | < 1 | 0.57 | 0.6
| Mean |  | 0.56667 | 0.535 | 0.49667 | 0.57482 | 0.61 |  | 0.58 | 0.62833
| Median |  | 0.6 | 0.53 | 0.495 | 0.58377 | 0.61 |  | 0.58 | 0.62
| Std Dev. |  | 0.05164 | 0.013784 | 0.017512 | 0.021394 | 0.016733 |  | 0.051769 | 0.036009
| Rel.Std.Dev. |  | 9.113% | 2.576% | 3.526% | 3.722% | 2.743% |  | 8.926% | 5.731%
| PDM3 |  | -0.6218% | -6.175% | -12.9% | 0.808% | 6.978% |  | 1.717% | 10.19%
| Z-Score (Absolute) |  | 0.07 | 0.64 | 1.33 | 0.07 | 0.71 | 1.27 | 0.17 | 1.04
| Analytical results for U in OREAS 553b (Certified Value 8.7 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 04 | 09 | 12 | 15 | 17 | 18 | 22 | 25 | 26 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [11 empty] | 2
3
| 1 | 1 | 9.5 | 8.9 | 8.4 | 8.6886 | 7.9 | 9.1512 | 9.44 | 8.6 | 8.5 | 8.63 | 9 | [11 empty] | 1
| 1 | 2 | 9.4 | 8.4 | 8.2 | 8.613 | 8.5 | 8.7378 | 9.03 | 9.4 | 8.5 | 8.19 | 9 | [11 empty] | 13
| 1 | 3 | 9.3 | 8.5 | 8.3 | 8.723 | 8.2 | 8.8003 | 9.06 | 8.8 | 8.3 | 8.4 | 9 | [11 empty] | 16
| 1 | 4 | 9.2 | 8.3 | 8.4 | 8.599 | 8.3 | 8.8442 | 9.25 | 8.8 | 8.5 | 8.7 | 9 | [11 empty] | 8.6995
| 1 | 5 | 9.3 | 9 | 8.4 | 8.68 | 7 | 9.0592 | 9.19 | 8.8 | 8.4 | 8.51 | 9 | [11 empty] | 141
| 1 | 6 | 9.4 | 7.7 | 8.4 | 8.7431 | 7.1 | 8.9061 | 9.39 | 8.8 | 8.4 | 7.16 | 9
| Mean |  | 9.35 | 8.4667 | 8.35 | 8.6745 | 7.8333 | 8.9165 | 9.2267 | 8.8667 | 8.4333 | 8.265 | 9
| Median |  | 9.35 | 8.45 | 8.4 | 8.6843 | 8.05 | 8.8752 | 9.22 | 8.8 | 8.45 | 8.455 | 9
| Std Dev. |  | 0.10488 | 0.46762 | 0.083666 | 0.057914 | 0.6377 | 0.15885 | 0.16765 | 0.27325 | 0.08165 | 0.5705 | 0
| Rel.Std.Dev. |  | 1.122% | 5.523% | 1.002% | 0.6676% | 8.141% | 1.782% | 1.817% | 3.082% | 0.9682% | 6.903% | 0%
| PDM3 |  | 7.477% | -2.677% | -4.018% | -0.2882% | -9.957% | 2.494% | 6.059% | 1.921% | -3.06% | -4.995% | 3.454%
| Z-Score (Absolute) |  | 1.75 | 0.23 | 0.49 | 0.23 | 1.65 | 0.78 | 1.47 | 0.66 | 0.31 | 0.68 | NA
Results from laboratory 26 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 553b (Certified Value 365 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 374 | 366 | 362 | 370 | 337.928 | 384 | 374 | 340 | 368.4014 | 366.8741 | 368 | 366 | 46.2 | 388 | 382 | [7 empty] | 1
| 1 | 2 | 374 | 357 | 373 | 370 | 331.5855 | 339 | 372 | 360 | 358.7645 | 367.8603 | 369 | 375 | 46.3 | 393 | 376 | [7 empty] | 29
| 1 | 3 | 368 | 350 | 396 | 360 | 335.1635 | 355 | 372 | 340 | 372.3669 | 365.1274 | 368 | 363 | 46.3 | 399 | 376 | [7 empty] | 16
| 1 | 4 | 367 | 357 | 386 | 360 | 331.582 | 340 | 373 | 360 | 359.601 | 363.6211 | 371 | 361 | 46.2 | 390 | 384 | [7 empty] | 364.947
| 1 | 5 | 364 | 358 | 391 | 360 | 335.237 | 324 | 373 | 350 | 357.2015 | 370.143 | 372 | 369 | 46.3 | 389 | 385 | [7 empty] | 142
| 1 | 6 | 372 | 356 | 374 | 360 | 330.0335 | 343 | 372 | 350 | 366.2463 | 366.8105 | 369 | 365 | 43.2 | 388 | 359
| Mean |  | 369.8333 | 357.3333 | 380.3333 | 363.3333 | 333.5883 | 347.5 | 372.6667 | 350 | 363.7636 | 366.7394 | 369.5 | 366.5 | 45.75 | 391.1667 | 377
| Median |  | 370 | 357 | 380 | 360 | 333.3745 | 341.5 | 372.5 | 350 | 362.9236 | 366.8423 | 369 | 365.5 | 46.25 | 389.5 | 379
| Std Dev. |  | 4.1191 | 5.1251 | 12.8167 | 5.164 | 2.9902 | 20.4426 | 0.8165 | 8.9443 | 6.1167 | 2.2444 | 1.6432 | 4.9699 | 1.2502 | 4.2622 | 9.6333
| Rel.Std.Dev. |  | 1.114% | 1.434% | 3.37% | 1.421% | 0.8964% | 5.883% | 0.2191% | 2.556% | 1.681% | 0.612% | 0.4447% | 1.356% | 2.733% | 1.09% | 2.555%
| PDM3 |  | 1.339% | -2.086% | 4.216% | -0.4422% | -8.593% | -4.781% | 2.115% | -4.096% | -0.3243% | 0.4911% | 1.248% | 0.4255% | -87.46% | 7.185% | 3.303%
| Z-Score (Absolute) |  | 0.25 | 0.67 | 1.02 | 0.23 | 2.42 | 1.4 | 0.45 | 1.21 | 0.2 | 0.02 | 0.22 | 0 | 23.59 | 1.81 | 0.77
| Analytical results for W in OREAS 553b (Certified Value 3.18 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 09 | 15 | 17 | 18 | 22 | 25 | 26 | [10 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | [10 empty] | 2
2
| 1 | 1 | 2 | < 100 | 3.3 | 4 | < 50 | 2.8439 | 3.0536 | 3 | 3 | < 5 | 3 | 4 | [10 empty] | 1
| 1 | 2 | 3 | < 100 | 3 | 4 | < 50 | 2.6548 | 2.9755 | 3 | 3 | < 5 | 3 | 3 | [10 empty] | 14
| 1 | 3 | 3 | < 100 | 2.9 | 4 | < 50 | 2.7061 | 2.9452 | 3 | 3 | < 5 | 3 | 3 | [10 empty] | 16
| 1 | 4 | 3 | < 100 | 3.1 | 4 | < 50 | 2.6262 | 3.0662 | 3 | 4 | < 5 | 3 | 4 | [10 empty] | 3.1755
| 1 | 5 | 3 | < 100 | 3.3 | 4 | < 50 | 2.919 | 3.0835 | 4 | 3 | < 5 | 3 | 4 | [10 empty] | 143
| 1 | 6 | 3 | < 100 | 2.9 | 4 | < 50 | 2.9177 | 3.1855 | 3 | 3 | < 5 | 2 | 4
| Mean |  | 2.8333 |  | 3.0833 | 4 |  | 2.7779 | 3.0516 | 3.1667 | 3.1667 |  | 2.8333 | 3.6667
| Median |  | 3 |  | 3.05 | 4 |  | 2.775 | 3.0599 | 3 | 3 |  | 3 | 4
| Std Dev. |  | 0.40825 |  | 0.18348 | 0 |  | 0.13199 | 0.085167 | 0.40825 | 0.40825 |  | 0.40825 | 0.5164
| Rel.Std.Dev. |  | 14.41% |  | 5.951% | 0% |  | 4.751% | 2.791% | 12.89% | 12.89% |  | 14.41% | 14.08%
| PDM3 |  | -10.78% |  | -2.902% | 25.96% |  | -12.52% | -3.903% | -0.2782% | -0.2782% |  | -10.78% | 15.47%
| Z-Score (Absolute) |  | 0.62 | 98.97 | 0.09 | 1.85 | 46.19 | 0.73 | 0.16 | 0.09 | 0.09 | 1.32 | 0.62 | 1.14
| Analytical results for Y in OREAS 553b (Certified Value 35.6 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | [7 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [7 empty] | 1
2
| 1 | 1 | 37.1 | 43 | 38.2 | 35.2 | 33.2 | 36 | 35 | 37.8079 | 34.3724 | 36.3 | 36.5 | 34 | 28.6 | 37 | 36.9 | [7 empty] | 1
| 1 | 2 | 35.7 | 40 | 39.1 | 35.1 | 32 | 36 | 36 | 35.6485 | 35.26 | 34.2 | 37.8 | 34 | 27.6 | 31.9 | 35.1 | [7 empty] | 15
| 1 | 3 | 36.2 | 41 | 37.6 | 34.4 | 33.1 | 36 | 34 | 37.3631 | 34.9642 | 34.1 | 36.4 | 37 | 28.1 | 43.5 | 37 | [7 empty] | 16
| 1 | 4 | 37.6 | 40 | 32.1 | 34.1 | 33.3 | 36 | 36 | 35.5248 | 35.6738 | 35.6 | 35.4 | 34 | 28.4 | 37.4 | 35.6 | [7 empty] | 35.6106
| 1 | 5 | 35.9 | 40 | 37.5 | 34.9 | 32.3 | 36 | 34 | 37.6484 | 34.5789 | 36.1 | 36.3 | 35 | 28.3 | 41.3 | 37.1 | [7 empty] | 144
| 1 | 6 | 38.1 | 42 | 29.8 | 34.7 | 32.4 | 36 | 35 | 37.9029 | 35.0634 | 36.2 | 35.2 | 34 | 28.5 | 40 | 35.3
| Mean |  | 36.7667 | 41 | 35.7167 | 34.7333 | 32.7167 | 36 | 35 | 36.9826 | 34.9855 | 35.4167 | 36.2667 | 34.6667 | 28.25 | 38.5167 | 36.1667
| Median |  | 36.65 | 40.5 | 37.55 | 34.8 | 32.75 | 36 | 35 | 37.5057 | 35.0138 | 35.85 | 36.35 | 34 | 28.35 | 38.7 | 36.25
| Std Dev. |  | 0.97912 | 1.2649 | 3.806 | 0.42269 | 0.54924 | 0 | 0.89443 | 1.0974 | 0.46836 | 1.0108 | 0.9288 | 1.2111 | 0.36194 | 4.0524 | 0.9288
| Rel.Std.Dev. |  | 2.663% | 3.085% | 10.66% | 1.217% | 1.679% | 0% | 2.556% | 2.967% | 1.339% | 2.854% | 2.561% | 3.493% | 1.281% | 10.52% | 2.568%
| PDM3 |  | 3.246% | 15.13% | 0.2978% | -2.464% | -8.127% | 1.093% | -1.715% | 3.853% | -1.756% | -0.5447% | 1.842% | -2.651% | -20.67% | 8.161% | 1.561%
| Z-Score (Absolute) |  | 0.72 | 3.62 | 0 | 0.67 | 2.06 | 0.19 | 0.49 | 0.87 | 0.5 | 0.21 | 0.38 | 0.72 | 5.12 | 1.92 | 0.31
| Analytical results for Yb in OREAS 553b (Certified Value 3.63 ppm) | [23 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 03 | 04 | 06 | 15 | 17 | 19 | 22 | 25 | 26 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | [13 empty] | 2
3
| 1 | 1 | 3.4 | 3.8 | 3.8 | 3.4485 | 4.2 | 3 | 3.2 | 3.3 | 4.1 | [13 empty] | 1
| 1 | 2 | 3.5 | 3.7 | 3.7 | 3.7308 | 3.8 | 3 | 3.1 | 3.6 | 4.2 | [13 empty] | 16
| 1 | 3 | 3.9 | 3.8 | 3.9 | 3.3049 | 3.9 | 4 | 3 | 3.5 | 4 | [13 empty] | 16
| 1 | 4 | 3.4 | 3.8 | 3.7 | 3.3326 | 3.9 | 4 | 3 | 3.3 | 4.3 | [13 empty] | 3.6337
| 1 | 5 | 3.8 | 3.8 | 3.7 | 3.6741 | 4 | 4 | 3 | 3.3 | 4 | [13 empty] | 145
| 1 | 6 | 3 | 3.8 | 3.7 | 3.6283 | 4 | 4 | 3.1 | 3.3 | 4
| Mean |  | 3.5 | 3.7833 | 3.75 | 3.5199 | 3.9667 | 3.6667 | 3.0667 | 3.3833 | 4.1
| Median |  | 3.45 | 3.8 | 3.7 | 3.5384 | 3.95 | 4 | 3.05 | 3.3 | 4.05
| Std Dev. |  | 0.32249 | 0.040825 | 0.083666 | 0.18244 | 0.13663 | 0.5164 | 0.08165 | 0.13292 | 0.12649
| Rel.Std.Dev. |  | 9.214% | 1.079% | 2.231% | 5.183% | 3.444% | 14.08% | 2.662% | 3.929% | 3.085%
| PDM3 |  | -3.68% | 4.117% | 3.2% | -3.133% | 9.162% | 0.9063% | -15.61% | -6.891% | 12.83%
| Z-Score (Absolute) |  | 0.45 | 0.5 | 0.39 | 0.39 | 1.12 | NA | 1.92 | 0.85 | 1.57
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 553b (Indicative Value 27.2 ppm) | [23 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 22 | 25 | 26 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS |  |  |  | 1
1
| 1 | 1 | 24.12 | 50 | 30 | < 20 | 12 | 20 | < 20 | < 20 | 40 | < 100 | 50 | < 20 | 17 | 57 | 33 | < 100 | 16 | 18 | 17 |  |  |  | 1
| 1 | 2 |  | 50 | < 30 | < 20 | 12 | 10 | < 20 | < 20 | 40 | < 100 | 50 | < 20 | 16 |  | 21 | < 100 | 16 | 18 | 17 |  |  |  | 17
| 1 | 3 |  | 50 | < 30 | < 20 | 12 | 10 | < 20 | < 20 | 40 | < 100 | 40 | < 20 | 14 |  | 21 | < 100 | 15.9 | 20 | 24 |  |  |  | 16
| 1 | 4 | 21.34 | < 50 | 30 | < 20 | 12 | 20 | < 20 | < 20 | 40 | < 100 | 40 | < 20 | 15 |  | 21 | < 100 | 15.9 | 19 | 20 |  |  |  | 27.2003
| 1 | 5 | 45.55 | 60 | 30 | < 20 | 12 | 20 | < 20 | < 20 | 40 | < 100 | 40 | < 20 | 16 | 20 | 24 | < 100 | 15.9 | 17 | 26 |  |  |  | 26
| 1 | 6 |  | < 50 | < 30 | < 20 | 12 | 10 | < 20 | 70 | 30 | < 100 | 50 | < 20 | 15 |  | 23 | < 100 | 15.9 | 19 | 16
| Mean |  | 30.3367 | 52.5 | 30 |  | 12 | 15 |  | 70 | 38.3333 |  | 45 |  | 15.5 | 38.5 | 23.8333 |  | 15.9333 | 18.5 | 20
| Median |  | 24.12 | 50 | 30 |  | 12 | 15 |  | 70 | 40 |  | 45 |  | 15.5 | 38.5 | 22 |  | 15.9 | 18.5 | 18.5
| Std Dev. |  | 13.2483 | 5 | 0 |  | 0 | 5.4772 |  |  | 4.0825 |  | 5.4772 |  | 1.0488 | 26.163 | 4.6655 |  | 0.05164 | 1.0488 | 4.1473
| Rel.Std.Dev. |  | 43.67% | 9.524% | 0% |  | 0% | 36.51% |  |  | 10.65% |  | 12.17% |  | 6.767% | 67.96% | 19.58% |  | 0.3241% | 5.669% | 20.74%
| PDM3 |  | 11.53% | 93.01% | 10.29% |  | -55.88% | -44.85% |  | 157.4% | 40.93% |  | 65.44% |  | -43.02% | 41.54% | -12.38% |  | -41.42% | -31.99% | -26.47%
| Z-Score (Absolute) |  | 0.7 | 1.57 | 0.17 | 0.67 | 0.54 | 0.34 | 0.67 | 0 | 1.24 | 2.02 | 1.69 | 0.67 | 0.3 | 1.25 | 0.26 | 2.02 | 0.27 | 0.1 | 0
| Analytical results for Zr in OREAS 553b (Indicative Value 241 ppm) | [23 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 15 | 17 | 19 | 22 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | [18 empty] | 0
0
| 1 | 1 | 252.3082 | 248 | 220 | 10.5 | [18 empty] | 1
| 1 | 2 | 247.4375 | 240 | 226 | 10.3 | [18 empty] | 21
| 1 | 3 | 250.5019 | 242 | 239 | 10.4 | [18 empty] | 16
| 1 | 4 | 251.2786 | 245 | 223 | 10.3 | [18 empty] | 241.1063
| 1 | 5 | 245.8951 | 249 | 236 | 10.4 | [18 empty] | 27
| 1 | 6 | 249.4913 | 251 | 224 | 10.4
| Mean |  | 249.4854 | 245.8333 | 228 | 10.3833
| Median |  | 249.9966 | 246.5 | 225 | 10.4
| Std Dev. |  | 2.4208 | 4.2622 | 7.6681 | 0.075277
| Rel.Std.Dev. |  | 0.9703% | 1.734% | 3.363% | 0.725%
| PDM3 |  | 3.475% | 1.961% | -5.436% | -95.69%
| Z-Score (Absolute) |  | 0.79 | 0.56 | 0.56 | 14.22
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C558: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C576: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C612: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C630: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C648: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C666: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C684: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C702: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C720: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C738: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C756: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C774: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C792: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C810: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C828: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C846: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C864: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C882: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1026: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 553b (Certified Value 0.077 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.08 | 0.1 | 0.04 | < 0.05 | 0.07 | 0.08 | < 0.05 | 0.09 | 0.08 | 0.09 | 0.07 | 0.09 | < 0.1 | < 0.1 | 0.07 | 0.08 | < 0.3 | < 0.1 | < 0.1 | 2.03 | < 1 | 0.07 | 0.08 | < 0.1 |  |  | 1
| 1 | 2 | 0.08 | 0.102 | 0.041 | < 0.05 | 0.07 | 0.07 | < 0.05 | 0.08 | 0.09 | 0.08 | 0.07 | 0.08 | < 0.1 | < 0.1 | 0.06 | 0.07 | < 0.3 | < 0.1 | < 0.1 | 2.05 | < 1 | 0.08 | 0.07 | < 0.1 |  |  | 4
| 1 | 3 | 0.08 | 0.082 | 0.05 | < 0.05 | 0.07 | 0.08 | < 0.05 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | < 0.1 | < 0.1 | 0.07 | 0.06 | < 0.3 | < 0.1 | < 0.1 | 2.04 | < 1 | 0.08 | 0.06 | < 0.1 |  |  | 16
| 1 | 4 | 0.09 | 0.085 | 0.042 | < 0.05 | 0.07 | 0.09 | < 0.05 | 0.08 | 0.08 | 0.09 | 0.07 | 0.08 | < 0.1 | < 0.1 | 0.07 | 0.07 | < 0.3 | < 0.1 | < 0.1 | 2.04 | < 1 | 0.07 | 0.07 | < 0.1 |  |  | 0.076744
| 1 | 5 | 0.08 | 0.076 | 0.05 | < 0.05 | 0.07 | 0.07 | < 0.05 | 0.08 | 0.08 | 0.07 | 0.07 | 0.07 | < 0.1 | < 0.1 | 0.07 | 0.07 | < 0.3 | < 0.1 | < 0.1 | 2.01 | < 1 | 0.07 | 0.08 | < 0.1 |  |  | 147
| 1 | 6 | 0.08 | 0.081 | 0.057 | < 0.05 | 0.06 | 0.09 | < 0.05 | 0.08 | 0.07 | 0.08 | 0.07 | 0.08 | < 0.1 | < 0.1 | 0.07 | 0.08 | < 0.3 | < 0.1 | < 0.1 | 2.09 | < 1 | 0.07 | 0.08 | < 0.1
| Mean |  | 0.081667 | 0.087667 | 0.046667 |  | 0.068333 | 0.08 |  | 0.081667 | 0.08 | 0.081667 | 0.07 | 0.08 |  |  | 0.068333 | 0.071667 |  |  |  | 2.0433 |  | 0.073333 | 0.073333
| Median |  | 0.08 | 0.0835 | 0.046 |  | 0.07 | 0.08 |  | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 |  |  | 0.07 | 0.07 |  |  |  | 2.04 |  | 0.07 | 0.075
| Std Dev. |  | 0.0040825 | 0.010746 | 0.0067429 |  | 0.0040825 | 0.0089443 |  | 0.0040825 | 0.0063246 | 0.0075277 | 0 | 0.0063246 |  |  | 0.0040825 | 0.0075277 |  |  |  | 0.026583 |  | 0.005164 | 0.008165
| Rel.Std.Dev. |  | 4.999% | 12.26% | 14.45% |  | 5.974% | 11.18% |  | 4.999% | 7.906% | 9.218% | 0% | 7.906% |  |  | 5.974% | 10.5% |  |  |  | 1.301% |  | 7.042% | 11.13%
| PDM3 |  | 6.415% | 14.23% | -39.19% |  | -10.96% | 4.243% |  | 6.415% | 4.243% | 6.415% | -8.787% | 4.243% |  |  | -10.96% | -6.615% |  |  |  | 2563% |  | -4.444% | -4.444%
| Z-Score (Absolute) |  | 0.51 | 0.84 | 1.43 | 2.63 | 0.23 | 0.42 | 2.63 | 0.51 | 0.42 | 0.51 | 0.14 | 0.42 | 1.25 | 1.25 | 0.23 | 0.05 | 4.3 | 1.25 | 1.25 | 109.23 | 23.69 | 0.05 | 0.05 | 1.25
| Analytical results for Al in OREAS 553b (Certified Value 0.94 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 1.1206 | 0.93 | 1.43 | 1.54 | 1.086 | 1.12 | 0.77 | 1.017 | 0.89 | 0.85 | 0.82 | 0.96 | 0.82 | 1.0275 | 0.8 | 0.97 | 1.9238 | 1.01 | 0.99 | 0.58 | 0.94495 | 0.979 | 2.36 | 0.885 |  |  | 1
| 1 | 2 | 1.1579 | 1.06 | 1.5 | 1.58 | 1.112 | 1.15 | 0.76 | 1.0299 | 0.9 | 0.87 | 0.83 | 0.92 | 0.82 | 1.0357 | 0.79 | 0.97 | 2.0132 | 0.95 | 0.98 | 0.58 | 0.9819 | 0.995 | 2.38 | 0.931 |  |  | #N/A
| 1 | 3 | 1.1179 | 1.06 | 1.2 | 1.56 | 1.096 | 1.12 | 0.78 | 0.97952 | 0.91 | 0.86 | 0.81 | 0.95 | 0.86 | 1.01 | 0.81 | 0.97 | 1.9587 | 0.96 | 0.97 | 0.58 | 0.94594 | 0.995 | 2.27 | 0.876 |  |  | 16
| 1 | 4 | 1.1202 | 0.95 | 1.51 | 1.58 | 1.115 | 1.11 | 0.77 | 1.0279 | 0.85 | 0.85 | 0.84 | 0.97 | 0.81 | 1.0119 | 0.79 | 1.01 | 2.0514 | 0.89 | 0.97 | 0.58 | 0.96325 | 0.931 | 2.78 | 0.815 |  |  | 0.93966
| 1 | 5 | 1.1583 | 0.86 | 1.17 | 1.58 | 1.077 | 1.13 | 0.76 | 0.98161 | 0.87 | 0.84 | 0.85 | 0.94 | 0.82 | 1.0229 | 0.82 | 0.98 | 2.0686 | 0.91 | 0.96 | 0.59 | 0.967 | 0.916 | 2.34 | 0.849 |  |  | 148
| 1 | 6 | 1.1377 | 0.93 | 1.15 | 1.53 | 1.092 | 1.12 | 0.78 | 1.0057 | 0.9 | 0.81 | 0.83 | 0.98 | 0.84 | 0.99259 | 0.81 | 0.99 | 2.0374 | 0.96 | 0.97 | 0.57 | 0.97231 | 1.016 | 2.23 | 0.921
| Mean |  | 1.1354 | 0.965 | 1.3267 | 1.5617 | 1.0963 | 1.125 | 0.77 | 1.0069 | 0.88667 | 0.84667 | 0.83 | 0.95333 | 0.82833 | 1.0168 | 0.80333 | 0.98167 | 2.0089 | 0.94667 | 0.97333 | 0.58 | 0.96256 | 0.972 | 2.3933 | 0.8795
| Median |  | 1.1292 | 0.94 | 1.315 | 1.57 | 1.094 | 1.12 | 0.77 | 1.0114 | 0.895 | 0.85 | 0.83 | 0.955 | 0.82 | 1.0174 | 0.805 | 0.975 | 2.0253 | 0.955 | 0.97 | 0.58 | 0.96512 | 0.987 | 2.35 | 0.8805
| Std Dev. |  | 0.018933 | 0.079687 | 0.17096 | 0.022286 | 0.014787 | 0.013784 | 0.0089443 | 0.022173 | 0.022509 | 0.020656 | 0.014142 | 0.021602 | 0.018348 | 0.015264 | 0.012111 | 0.016021 | 0.056503 | 0.042269 | 0.010328 | 0.0063246 | 0.014667 | 0.039648 | 0.19775 | 0.043606
| Rel.Std.Dev. |  | 1.667% | 8.258% | 12.89% | 1.427% | 1.349% | 1.225% | 1.162% | 2.202% | 2.539% | 2.44% | 1.704% | 2.266% | 2.215% | 1.501% | 1.508% | 1.632% | 2.813% | 4.465% | 1.061% | 1.09% | 1.524% | 4.079% | 8.263% | 4.958%
| PDM3 |  | 20.83% | 2.697% | 41.19% | 66.19% | 16.67% | 19.72% | -18.06% | 7.158% | -5.64% | -9.897% | -11.67% | 1.455% | -11.85% | 8.206% | -14.51% | 4.47% | 113.8% | 0.7455% | 3.583% | -38.28% | 2.437% | 3.442% | 154.7% | -6.402%
| Z-Score (Absolute) |  | 0.9 | 0.02 | 1.93 | 3.2 | 0.69 | 0.85 | 1.07 | 0.21 | 0.44 | 0.66 | 0.75 | 0.08 | 0.76 | 0.26 | 0.89 | 0.07 | 5.62 | 0.12 | 0.03 | 2.1 | 0.03 | 0.02 | 7.7 | 0.48
| Analytical results for As in OREAS 553b (Certified Value 51 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 1
| 1 | 1 | 55 | 47.7 | 49.5 | 51 | 47.99 | 57 | 51 | 53 | 53.5352 | 54.4 | 51.1 | 54.1 | 57.3 | 55.3 | 51.323 | 47.6 | 49 | 53 | 51.7 | 52.4 | 51.6 | 40.9 | 42.25 | 35.6 | 48.5 | 52 | 1
| 1 | 2 | 54 | 48 | 51.3 | 52 | 48.19 | 60 | 54 | 54 | 52.7951 | 54.8 | 52.1 | 54.9 | 53.8 | 54.6 | 52.7181 | 46.1 | 51 | 55 | 54.1 | 48.1 | 52.6 | 40.5 | 44.285 | 36 | 47.1 | 51 | 5
| 1 | 3 | 55 | 47.5 | 48.7 | 53 | 50.19 | 57 | 51 | 51 | 53.2638 | 54.2 | 52.7 | 51.2 | 55.6 | 52.1 | 52.9487 | 48.6 | 50 | 54 | 51.2 | 50.4 | 51.8 | 41.1 | 42.369 | 37.1 | 48.9 | 49 | 16
| 1 | 4 | 54 | 47 | 53 | 51 | 47.23 | 62 | 54 | 52 | 52.5559 | 51 | 51.4 | 53.4 | 55.4 | 53.2 | 51.0941 | 49.2 | 50 | 55 | 52.2 | 49.8 | 50.1 | 39 | 43.489 | 37 | 49.1 | 47 | 51.3639
| 1 | 5 | 55 | 46 | 48.6 | 51 | 50.73 | 58 | 53 | 53 | 53.0922 | 50.9 | 49.3 | 52.5 | 56.4 | 52.1 | 50.1692 | 48.2 | 50 | 55 | 49 | 49.9 | 50.5 | 40.7 | 42.853 | 35.6 | 48 | 49 | 149
| 1 | 6 | 55 | 45.6 | 49.6 | 51 | 48.5 | 62 | 51 | 54 | 53.0339 | 52.3 | 49.7 | 52.8 | 55 | 50.6 | 50.3723 | 46.3 | 50 | 54 | 51.5 | 50.1 | 50.1 | 40 | 43.571 | 38 | 46.9 | 50
| Mean |  | 54.6667 | 46.9667 | 50.1167 | 51.5 | 48.805 | 59.3333 | 52.3333 | 52.8333 | 53.046 | 52.9333 | 51.05 | 53.15 | 55.5833 | 52.9833 | 51.4375 | 47.6667 | 50 | 54.3333 | 51.6167 | 50.1167 | 51.1167 | 40.3667 | 43.1362 | 36.55 | 48.0833 | 49.6667
| Median |  | 55 | 47.25 | 49.55 | 51 | 48.345 | 59 | 52 | 53 | 53.0631 | 53.25 | 51.25 | 53.1 | 55.5 | 52.65 | 51.2085 | 47.9 | 50 | 54.5 | 51.6 | 50 | 51.05 | 40.6 | 43.171 | 36.5 | 48.25 | 49.5
| Std Dev. |  | 0.5164 | 0.96885 | 1.7128 | 0.83666 | 1.3594 | 2.3381 | 1.5055 | 1.169 | 0.34414 | 1.7614 | 1.3293 | 1.2942 | 1.1974 | 1.7475 | 1.166 | 1.2517 | 0.63246 | 0.8165 | 1.6461 | 1.379 | 1.0342 | 0.76855 | 0.7858 | 0.97108 | 0.92177 | 1.7512
| Rel.Std.Dev. |  | 0.9446% | 2.063% | 3.418% | 1.625% | 2.785% | 3.941% | 2.877% | 2.213% | 0.6488% | 3.328% | 2.604% | 2.435% | 2.154% | 3.298% | 2.267% | 2.626% | 1.265% | 1.503% | 3.189% | 2.752% | 2.023% | 1.904% | 1.822% | 2.657% | 1.917% | 3.526%
| PDM3 |  | 6.43% | -8.561% | -2.428% | 0.265% | -4.982% | 15.52% | 1.887% | 2.861% | 3.275% | 3.056% | -0.6111% | 3.477% | 8.215% | 3.153% | 0.1434% | -7.198% | -2.655% | 5.781% | 0.4922% | -2.428% | -0.4813% | -21.41% | -16.02% | -28.84% | -6.387% | -3.304%
| Z-Score (Absolute) |  | 1.32 | 1.67 | 0.45 | 0.09 | 0.96 | 3.13 | 0.41 | 0.6 | 0.69 | 0.64 | 0.09 | 0.73 | 1.67 | 0.66 | 0.06 | 1.4 | 0.5 | 1.19 | 0.13 | 0.45 | 0.06 | 4.23 | 3.16 | 5.72 | 1.24 | 0.62
| Analytical results for Au in OREAS 553b (Certified Value < 0.02 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 02 | 03 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 20 | 24 | 25 | [13 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [13 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.1g | [13 empty] | 3
| 1 | 1 |  | 0.0013 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 |  | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | 1
| 1 | 2 |  | 0.0026 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 |  | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | #N/A
| 1 | 3 |  | 0.0002 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 |  | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | 16
| 1 | 4 | 0.0001 | 0.0029 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 |  | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | < 0.02
| 1 | 5 |  | 0.0028 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 |  | < 0.0005 | < 0.01 | < 0.0005 | [13 empty] | 150
| 1 | 6 |  | 0.0041 | < 0.0005 | < 0.001 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.003 | < 0.0005 | < 0.01 | < 0.0005
| Mean |  | 0.0001 | 0.0023167 | [7 empty] | 0.003
| Median |  | 0.0001 | 0.0027 | [7 empty] | 0.003
| Std Dev. |  |  | 0.0013674
| Rel.Std.Dev. |  |  | 59.02%
PDM3
| Z-Score (Absolute) |  | 0.71 | 0.17 | 0.67 | 0.61 | 1.72 | 1.72 | 1.72 | 1.72 | 1.72 | 0 | 0.67 | 0.49 | 0.67
| Analytical results for B in OREAS 553b (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 03 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 18 | 20 | 25 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 01g | 0.5g | 0.1g | [12 empty] | 1
| 1 | 1 | < 20 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3.9295 |  | < 20 | 16 | [12 empty] | 1
| 1 | 2 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3.5466 |  | < 20 | 16 | [12 empty] | 1
| 1 | 3 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.1742 |  | < 20 | 17 | [12 empty] | 16
| 1 | 4 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3.9666 |  | < 20 | 16 | [12 empty] | < 10
| 1 | 5 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 3.5769 | 11 | < 20 | 15 | [12 empty] | 151
| 1 | 6 | < 20 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 4.0793 |  | < 20 | 19
| Mean |  |  | 4.1667 | [8 empty] | 3.8789 | 11 |  | 16.5
| Median |  |  | 4 | [8 empty] | 3.9481 | 11 |  | 16
| Std Dev. |  |  | 0.40825 | [8 empty] | 0.2605 |  |  | 1.3784
| Rel.Std.Dev. |  |  | 9.798% | [8 empty] | 6.716% |  |  | 8.354%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 553b (Certified Value 49.9 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 1
| 1 | 1 | 53 | 48.7 | 10.7 | 66 | 62 | 38 | 50 | 61.1159 | 50 | 50 | 50 | 50 | 50 | 53.8021 | 43.7 | 50 | 72 | 49 | 58 | 46.3 | 46.547 | 91 | 86 | 41 |  |  | 1
| 1 | 2 | 52 | 51.8 | 9.6 | 66 | 62 | 39 | 47 | 60.1379 | 50 | 50 | 50 | 50 | 50 | 56.2149 | 41.7 | 51 | 75 | 45 | 60 | 46.6 | 48.491 | 131 | 86.8 | 43 |  |  | #N/A
| 1 | 3 | 52 | 52.3 | 12.4 | 66 | 63 | 38 | 49 | 64.2355 | 50 | 50 | 50 | 50 | 50 | 55.2636 | 43.3 | 51 | 71 | 46 | 58 | 48 | 46.765 | 52 | 80.4 | 41 |  |  | 16
| 1 | 4 | 52 | 49 | 9.6 | 66 | 61 | 38 | 48 | 63.0713 | 50 | 50 | 50 | 50 | 50 | 53.1788 | 43 | 52 | 73 | 45 | 51 | 47.1 | 47.196 | 77 | 90.2 | 40 |  |  | 49.8917
| 1 | 5 | 52 | 45.9 | 8.8 | 64 | 66 | 38 | 49 | 60.734 | 50 | 50 | 50 | 50 | 50 | 53.346 | 42.3 | 50 | 74 | 45 | 50 | 48.9 | 47.851 | 79 | 82.7 | 40 |  |  | 152
| 1 | 6 | 53 | 48.7 | 12.3 | 65 | 64 | 38 | 48 | 62.6834 | 50 | 50 | 50 | 50 | 50 | 53.3893 | 42 | 51 | 73 | 48 | 50 | 46.7 | 47.63 | 64 | 82.5 | 43
| Mean |  | 52.3333 | 49.4 | 10.5667 | 65.5 | 63 | 38.1667 | 48.5 | 61.9963 | 50 | 50 | 50 | 50 | 50 | 54.1991 | 42.6667 | 50.8333 | 73 | 46.3333 | 54.5 | 47.2667 | 47.4133 | 82.3333 | 84.7667 | 41.3333
| Median |  | 52 | 48.85 | 10.15 | 66 | 62.5 | 38 | 48.5 | 61.8996 | 50 | 50 | 50 | 50 | 50 | 53.5957 | 42.65 | 51 | 73 | 45.5 | 54.5 | 46.9 | 47.413 | 78 | 84.35 | 41
| Std Dev. |  | 0.5164 | 2.3478 | 1.5082 | 0.83666 | 1.7889 | 0.40825 | 1.0488 | 1.5789 | 0 | 0 | 0 | 0 | 0 | 1.2473 | 0.78655 | 0.75277 | 1.4142 | 1.7512 | 4.6368 | 0.99331 | 0.7236 | 27.3472 | 3.568 | 1.3663
| Rel.Std.Dev. |  | 0.9867% | 4.753% | 14.27% | 1.277% | 2.839% | 1.07% | 2.162% | 2.547% | 0% | 0% | 0% | 0% | 0% | 2.301% | 1.843% | 1.481% | 1.937% | 3.78% | 8.508% | 2.102% | 1.526% | 33.22% | 4.209% | 3.305%
| PDM3 |  | 4.894% | -0.9855% | -78.82% | 31.28% | 26.27% | -23.5% | -2.789% | 24.26% | 0.2171% | 0.2171% | 0.2171% | 0.2171% | 0.2171% | 8.634% | -14.48% | 1.887% | 46.32% | -7.132% | 9.237% | -5.261% | -4.967% | 65.02% | 69.9% | -17.15%
| Z-Score (Absolute) |  | 0.4 | 0.1 | 6.76 | 2.66 | 2.23 | 2.03 | 0.26 | 2.06 | 0 | 0 | 0 | 0 | 0 | 0.72 | 1.26 | 0.14 | 3.94 | 0.63 | 0.77 | 0.47 | 0.44 | 5.54 | 5.96 | 1.49
| Analytical results for Be in OREAS 553b (Certified Value 2.16 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 3
| 1 | 1 | 2.18 | 2.4 | 2.58 | 5.6123 | 2.4 | 2.05 | 2.1 | 2.0639 | 2.12 | 2.06 | 2 | 2.23 | 1.99 | 2.1645 | 2 | 2.2 | 2 | 2.2 | 2.318 | 1.91 | 2.16 | 2.4 |  |  |  |  | 1
| 1 | 2 | 2.21 | 2.4 | 2.58 | 5.5913 | 2.4 | 1.96 | 1.9 | 2.0731 | 2.09 | 2.08 | 1.98 | 2.19 | 1.97 | 2.1882 | 1.9 | 2.2 | 2 | 2.2 | 2.396 | 1.96 | 1.91 | 2.5 |  |  |  |  | 6
| 1 | 3 | 2.23 | 2 | 2.58 | 5.5919 | 2.4 | 1.92 | 2 | 2.194 | 2.11 | 2.04 | 1.98 | 2.19 | 2.01 | 2.1318 | 1.8 | 2.2 | 2 | 2.2 | 2.304 | 2 | 2.13 | 2.4 |  |  |  |  | 16
| 1 | 4 | 2.23 | 2.4 | 2.55 | 5.6132 | 2.5 | 1.89 | 2.2 | 2.2433 | 2.04 | 2.04 | 2 | 2.23 | 1.95 | 2.1295 | 2.1 | 2.3 | 2 | 2.2 | 2.29 | 2 | 2.19 | 2.3 |  |  |  |  | 2.1563
| 1 | 5 | 2.22 | 1.9 | 2.56 | 5.5989 | 2.5 | 1.87 | 1.9 | 2.0385 | 2.05 | 2.02 | 2.06 | 2.19 | 1.94 | 2.1782 | 1.9 | 2.3 | 2 | 2.2 | 2.454 | 1.93 | 2.14 | 2.4 |  |  |  |  | 153
| 1 | 6 | 2.19 | 2 | 2.59 | 5.6218 | 2.5 | 2 | 2.1 | 2.1599 | 2.09 | 2 | 2.05 | 2.23 | 1.96 | 2.1262 | 1.9 | 2.3 | 2 | 2.2 | 1.89 | 1.99 | 2.11 | 2.4
| Mean |  | 2.21 | 2.1833 | 2.5733 | 5.6049 | 2.45 | 1.9483 | 2.0333 | 2.1288 | 2.0833 | 2.04 | 2.0117 | 2.21 | 1.97 | 2.1531 | 1.9333 | 2.25 | 2 | 2.2 | 2.2753 | 1.965 | 2.1067 | 2.4
| Median |  | 2.215 | 2.2 | 2.58 | 5.6056 | 2.45 | 1.94 | 2.05 | 2.1165 | 2.09 | 2.04 | 2 | 2.21 | 1.965 | 2.1481 | 1.9 | 2.25 | 2 | 2.2 | 2.311 | 1.975 | 2.135 | 2.4
| Std Dev. |  | 0.020976 | 0.24014 | 0.015055 | 0.012643 | 0.054772 | 0.068532 | 0.12111 | 0.082217 | 0.032042 | 0.028284 | 0.034881 | 0.021909 | 0.026077 | 0.027298 | 0.10328 | 0.054772 | 0 | 4.8648e-16 | 0.19891 | 0.038341 | 0.10013 | 0.063246
| Rel.Std.Dev. |  | 0.9491% | 11% | 0.5851% | 0.2256% | 2.236% | 3.517% | 5.956% | 3.862% | 1.538% | 1.386% | 1.734% | 0.9914% | 1.324% | 1.268% | 5.342% | 2.434% | 0% | 2.211e-14% | 8.742% | 1.951% | 4.753% | 2.635%
| PDM3 |  | 2.492% | 1.255% | 19.34% | 159.9% | 13.62% | -9.644% | -5.702% | -1.275% | -3.383% | -5.392% | -6.706% | 2.492% | -8.639% | -0.1491% | -10.34% | 4.347% | -7.247% | 2.028% | 5.521% | -8.871% | -2.301% | 11.3%
| Z-Score (Absolute) |  | 0.32 | 0.17 | 2.37 | 19.44 | 1.67 | 1.15 | 0.67 | 0.14 | 0.39 | 0.64 | 0.8 | 0.32 | 1.03 | 0 | 1.24 | 0.55 | NA | 0.26 | 0.69 | 1.06 | 0.26 | 1.39
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 553b (Certified Value 1.58 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 1.69 | 1.58 | 1.38 | 1.66 | 1.6 | 1.57 | 1.4206 | 1.63 | 1.56 | 1.62 | 1.52 | 1.78 | 1.6016 | 1.64 | 1.59 | 1.74 | 2.33 | 1.7 | 1.4 | < 5 | 1.418 | 1.51 | 1.78 | 1.55 |  |  | 1
| 1 | 2 | 1.67 | 1.64 | 1.39 | 1.65 | 1.63 | 1.6 | 1.3813 | 1.63 | 1.55 | 1.63 | 1.44 | 1.7 | 1.7273 | 1.61 | 1.59 | 1.79 | 2.27 | 1.5 | 1.4 | < 5 | 1.504 | 1.53 | 1.74 | 1.59 |  |  | 7
| 1 | 3 | 1.67 | 1.6 | 1.33 | 1.65 | 1.6 | 1.58 | 1.3986 | 1.63 | 1.55 | 1.58 | 1.5 | 1.72 | 1.7366 | 1.65 | 1.58 | 1.73 | 2.4 | 1.6 | 1.4 | < 5 | 1.447 | 1.56 | 1.72 | 1.52 |  |  | 16
| 1 | 4 | 1.76 | 1.6 | 1.46 | 1.65 | 1.61 | 1.57 | 1.4421 | 1.54 | 1.54 | 1.6 | 1.46 | 1.7 | 1.7333 | 1.61 | 1.58 | 1.76 | 2.39 | 1.6 | 1.4 | < 5 | 1.494 | 1.55 | 1.76 | 1.56 |  |  | 1.5791
| 1 | 5 | 1.69 | 1.59 | 1.29 | 1.66 | 1.57 | 1.62 | 1.4483 | 1.5 | 1.52 | 1.56 | 1.48 | 1.72 | 1.6228 | 1.59 | 1.56 | 1.79 | 2.33 | 1.6 | 1.4 | < 5 | 1.416 | 1.5 | 1.71 | 1.54 |  |  | 154
| 1 | 6 | 1.75 | 1.58 | 1.35 | 1.65 | 1.61 | 1.56 | 1.3866 | 1.59 | 1.52 | 1.6 | 1.51 | 1.69 | 1.6239 | 1.62 | 1.53 | 1.77 | 2.37 | 1.6 | 1.4 | < 5 | 1.461 | 1.53 | 1.69 | 1.48
| Mean |  | 1.705 | 1.5983 | 1.3667 | 1.6533 | 1.6033 | 1.5833 | 1.4129 | 1.5867 | 1.54 | 1.5983 | 1.485 | 1.7183 | 1.6743 | 1.62 | 1.5717 | 1.7633 | 2.3483 | 1.6 | 1.4 |  | 1.4567 | 1.53 | 1.7333 | 1.54
| Median |  | 1.69 | 1.595 | 1.365 | 1.65 | 1.605 | 1.575 | 1.4096 | 1.61 | 1.545 | 1.6 | 1.49 | 1.71 | 1.6756 | 1.615 | 1.58 | 1.765 | 2.35 | 1.6 | 1.4 |  | 1.454 | 1.53 | 1.73 | 1.545
| Std Dev. |  | 0.039875 | 0.022286 | 0.058195 | 0.005164 | 0.019664 | 0.022509 | 0.028502 | 0.055377 | 0.016733 | 0.025626 | 0.030822 | 0.032506 | 0.064276 | 0.021909 | 0.023166 | 0.025033 | 0.048339 | 0.063246 | 2.4324e-16 |  | 0.037136 | 0.022804 | 0.033267 | 0.037417
| Rel.Std.Dev. |  | 2.339% | 1.394% | 4.258% | 0.3123% | 1.226% | 1.422% | 2.017% | 3.49% | 1.087% | 1.603% | 2.076% | 1.892% | 3.839% | 1.352% | 1.474% | 1.42% | 2.058% | 3.953% | 1.737e-14% |  | 2.549% | 1.49% | 1.919% | 2.43%
| PDM3 |  | 7.972% | 1.217% | -13.45% | 4.7% | 1.534% | 0.2672% | -10.52% | 0.4783% | -2.477% | 1.217% | -5.96% | 8.816% | 6.025% | 2.589% | -0.4716% | 11.67% | 48.71% | 1.323% | -11.34% |  | -7.754% | -3.11% | 9.766% | -2.477%
| Z-Score (Absolute) |  | 1 | 0 | 2.17 | 0.51 | 0.05 | 0.14 | 1.73 | 0.11 | 0.55 | 0 | 1.06 | 1.12 | 0.71 | 0.2 | 0.25 | 1.54 | 7.01 | 0.02 | 1.85 | 8.43 | 1.32 | 0.64 | 1.26 | 0.55
| Analytical results for Ca in OREAS 553b (Certified Value 0.3 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 25 | 26 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 3
| 1 | 1 | 0.3 | 0.29 | 0.29 | 0.29 | 0.27689 | 0.313 | 0.3 | 0.3 | 0.30763 | 0.32 | 0.32 | 0.31 | 0.29 | 0.3 | 0.29906 | 0.3 | 0.29 | 0.31 | 0.3009 | 0.31 | 0.27 | 0.29 | 0.322 | 0.283 | 0.307 |  | 1
| 1 | 2 | 0.32 | 0.3 | 0.3 | 0.29 | 0.27535 | 0.309 | 0.3 | 0.29 | 0.30599 | 0.31 | 0.32 | 0.3 | 0.29 | 0.3 | 0.30034 | 0.29 | 0.29 | 0.31 | 0.2999 | 0.29 | 0.26 | 0.3 | 0.329 | 0.258 | 0.313 |  | #N/A
| 1 | 3 | 0.31 | 0.29 | 0.28 | 0.3 | 0.27596 | 0.317 | 0.3 | 0.31 | 0.30481 | 0.31 | 0.32 | 0.3 | 0.29 | 0.31 | 0.29255 | 0.3 | 0.29 | 0.32 | 0.3021 | 0.3 | 0.26 | 0.29 | 0.336 | 0.261 | 0.305 |  | 16
| 1 | 4 | 0.31 | 0.29 | 0.3 | 0.3 | 0.27667 | 0.318 | 0.3 | 0.29 | 0.29923 | 0.31 | 0.32 | 0.3 | 0.29 | 0.3 | 0.29413 | 0.29 | 0.3 | 0.31 | 0.3085 | 0.31 | 0.26 | 0.29 | 0.329 | 0.262 | 0.304 |  | 0.30016
| 1 | 5 | 0.32 | 0.29 | 0.28 | 0.29 | 0.27618 | 0.321 | 0.3 | 0.31 | 0.29961 | 0.3 | 0.31 | 0.31 | 0.29 | 0.3 | 0.29947 | 0.3 | 0.29 | 0.32 | 0.2892 | 0.3 | 0.26 | 0.3 | 0.322 | 0.285 | 0.306 |  | 155
| 1 | 6 | 0.31 | 0.28 | 0.27 | 0.29 | 0.27608 | 0.325 | 0.3 | 0.3 | 0.30731 | 0.31 | 0.31 | 0.32 | 0.29 | 0.3 | 0.29444 | 0.29 | 0.29 | 0.31 | 0.3007 | 0.3 | 0.26 | 0.29 | 0.314 | 0.284 | 0.311
| Mean |  | 0.31167 | 0.29 | 0.28667 | 0.29333 | 0.27619 | 0.31717 | 0.3 | 0.3 | 0.3041 | 0.31 | 0.31667 | 0.30667 | 0.29 | 0.30167 | 0.29666 | 0.295 | 0.29167 | 0.31333 | 0.30022 | 0.30167 | 0.26167 | 0.29333 | 0.32533 | 0.27217 | 0.30767
| Median |  | 0.31 | 0.29 | 0.285 | 0.29 | 0.27613 | 0.3175 | 0.3 | 0.3 | 0.3054 | 0.31 | 0.32 | 0.305 | 0.29 | 0.3 | 0.29675 | 0.295 | 0.29 | 0.31 | 0.3008 | 0.3 | 0.26 | 0.29 | 0.3255 | 0.2725 | 0.3065
| Std Dev. |  | 0.0075277 | 0.0063246 | 0.012111 | 0.005164 | 0.00054609 | 0.0056716 | 0 | 0.0089443 | 0.0037606 | 0.0063246 | 0.005164 | 0.008165 | 0 | 0.0040825 | 0.0033297 | 0.0054772 | 0.0040825 | 0.005164 | 0.0062342 | 0.0075277 | 0.0040825 | 0.005164 | 0.0076333 | 0.013045 | 0.003559
| Rel.Std.Dev. |  | 2.415% | 2.181% | 4.225% | 1.76% | 0.1977% | 1.788% | 0% | 2.981% | 1.237% | 2.04% | 1.631% | 2.662% | 0% | 1.353% | 1.122% | 1.857% | 1.4% | 1.648% | 2.077% | 2.495% | 1.56% | 1.76% | 2.346% | 4.793% | 1.157%
| PDM3 |  | 3.833% | -3.385% | -4.496% | -2.275% | -7.986% | 5.666% | -0.05347% | -0.05347% | 1.312% | 3.278% | 5.499% | 2.168% | -3.385% | 0.5018% | -1.165% | -1.719% | -2.83% | 4.389% | 0.01871% | 0.5018% | -12.82% | -2.275% | 8.386% | -9.326% | 2.501%
| Z-Score (Absolute) |  | 0.94 | 0.81 | 1.08 | 0.54 | 1.93 | 1.39 | 0 | 0 | 0.33 | 0.81 | 1.35 | 0.54 | 0.81 | 0.13 | 0.27 | 0.4 | 0.67 | 1.08 | 0.02 | 0.13 | 3.1 | 0.54 | 2.04 | 2.25 | 0.62
| Analytical results for Cd in OREAS 553b (Certified Value 0.3 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.33 | 0.31 | 0.24 | 0.31 | < 1 | 0.3 | 0.33 | 0.27847 | 0.3 | 0.29 | 0.27 | 0.26 | 0.3 | 0.32474 | 0.28 | 0.27 |  | 0.3 | 0.3 | < 1 | 0.272 | 0.29 | 0.18 | 0.31 |  |  | 1
| 1 | 2 | 0.31 | 0.33 | 0.28 | 0.32 | 1 | 0.36 | 0.34 | 0.29058 | 0.3 | 0.29 | 0.29 | 0.25 | 0.29 | 0.32365 | 0.29 | 0.27 |  | 0.3 | 0.3 | < 1 | 0.287 | 0.31 | 0.18 | 0.31 |  |  | 8
| 1 | 3 | 0.26 | 0.35 | 0.26 | 0.32 | 1 | 0.35 | 0.32 | 0.31372 | 0.28 | 0.31 | 0.27 | 0.25 | 0.3 | 0.30899 | 0.31 | 0.31 |  | 0.3 | 0.3 | < 1 | 0.284 | 0.31 | 0.18 | 0.31 |  |  | 16
| 1 | 4 | 0.31 | 0.33 | 0.3 | 0.3 | 1 | 0.34 | 0.33 | 0.27821 | 0.28 | 0.28 | 0.28 | 0.26 | 0.29 | 0.33399 | 0.29 | 0.27 | 0.32 | 0.3 | 0.3 | < 1 | 0.27 | 0.32 | 0.19 | 0.31 |  |  | 0.29845
| 1 | 5 | 0.31 | 0.31 | 0.25 | 0.3 | 1 | 0.32 | 0.31 | 0.31645 | 0.3 | 0.27 | 0.29 | 0.26 | 0.28 | 0.29605 | 0.34 | 0.29 | 0.31 | 0.3 | 0.3 | < 1 | 0.238 | 0.3 | 0.18 | 0.3 |  |  | 156
| 1 | 6 | 0.33 | 0.26 | 0.27 | 0.31 | 1 | 0.32 | 0.32 | 0.3011 | 0.31 | 0.29 | 0.3 | 0.26 | 0.29 | 0.29746 | 0.31 | 0.28 |  | 0.3 | 0.3 | < 1 | 0.29 | 0.31 | 0.19 | 0.29
| Mean |  | 0.30833 | 0.315 | 0.26667 | 0.31 | 1 | 0.33167 | 0.325 | 0.29642 | 0.295 | 0.28833 | 0.28333 | 0.25667 | 0.29167 | 0.31414 | 0.30333 | 0.28167 | 0.315 | 0.3 | 0.3 |  | 0.2735 | 0.30667 | 0.18333 | 0.305
| Median |  | 0.31 | 0.32 | 0.265 | 0.31 | 1 | 0.33 | 0.325 | 0.29584 | 0.3 | 0.29 | 0.285 | 0.26 | 0.29 | 0.31632 | 0.3 | 0.275 | 0.315 | 0.3 | 0.3 |  | 0.278 | 0.31 | 0.18 | 0.31
| Std Dev. |  | 0.025626 | 0.030822 | 0.021602 | 0.0089443 | 0 | 0.022286 | 0.010488 | 0.016796 | 0.012247 | 0.013292 | 0.012111 | 0.005164 | 0.0075277 | 0.015676 | 0.021602 | 0.016021 | 0.0070711 | 0 | 0 |  | 0.019181 | 0.010328 | 0.005164 | 0.0083666
| Rel.Std.Dev. |  | 8.311% | 9.785% | 8.101% | 2.885% | 0% | 6.719% | 3.227% | 5.666% | 4.152% | 4.61% | 4.274% | 2.012% | 2.581% | 4.99% | 7.122% | 5.688% | 2.245% | 0% | 0% |  | 7.013% | 3.368% | 2.817% | 2.743%
| PDM3 |  | 3.312% | 5.546% | -10.65% | 3.871% | 235.1% | 11.13% | 8.897% | -0.6791% | -1.155% | -3.389% | -5.064% | -14% | -2.272% | 5.259% | 1.637% | -5.623% | 5.546% | 0.5202% | 0.5202% |  | -8.359% | 2.754% | -38.57% | 2.196%
| Z-Score (Absolute) |  | 0.34 | 0.67 | 1.77 | 0.42 | 31.1 | 1.52 | 1.18 | 0.27 | 0.34 | 0.67 | 0.93 | 2.28 | 0.51 | 0.63 | 0.08 | 1.01 | 0.67 | 0.08 | 0.08 | 10.03 | 1.42 | 0.25 | 5.98 | 0.17
| Analytical results for Ce in OREAS 553b (Certified Value 42.9 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 1
| 1 | 1 | 47.988 | 54.8 | 57.9 | 85.72 | 42.46 | 32.36 | 43.1242 | 44.1 | 42.2 | 39.4 | 47.2 | 42.8 | 45.9719 | 34.2 | 39.89 | 71.11 | 42 | 36.9 | 45.011 | 90.8 | 101 | 36.15 |  |  |  |  | 1
| 1 | 2 | 47.007 | 56 | 58.5 | 85.664 | 44.04 | 31.7 | 45.5793 | 43.8 | 43.5 | 39.9 | 45 | 43 | 48.0296 | 33.1 | 40.57 | 73.51 | 41 | 38.5 | 46.014 | 75.4 | 92.1 | 37.34 |  |  |  |  | 9
| 1 | 3 | 47.189 | 51.5 | 58 | 85.732 | 42.78 | 33.55 | 43.4822 | 43.1 | 43.3 | 38.6 | 46.1 | 43.8 | 45.9647 | 34 | 42.03 | 70.89 | 41 | 36.1 | 44.982 | 77.4 | 105 | 35.65 |  |  |  |  | 16
| 1 | 4 | 49.376 | 53.4 | 58.5 | 85.696 | 42.14 | 33.83 | 44.7832 | 39.7 | 42.5 | 40 | 47 | 42.8 | 45.4476 | 34 | 42.48 | 72.94 | 41 | 35.8 | 44.634 | 72.6 | 104 | 34.34 |  |  |  |  | 42.8634
| 1 | 5 | 47.883 | 49.4 | 57.1 | 85.684 | 41.88 | 32.07 | 40.9419 | 42.1 | 41.8 | 38.6 | 46.1 | 40.7 | 45.697 | 33.5 | 40.81 | 72.8 | 43 | 38.5 | 45.055 | 84 | 94.2 | 34.43 |  |  |  |  | 157
| 1 | 6 | 49.98 | 52.5 | 58.4 | 85.712 | 41.74 | 33.23 | 41.6023 | 41.9 | 40.7 | 40.3 | 45.3 | 41.5 | 46.0316 | 33.8 | 41.82 | 73.87 | 42 | 36.9 | 45.671 | 76.1 | 96.2 | 36.41
| Mean |  | 48.2372 | 52.9333 | 58.0667 | 85.7013 | 42.5067 | 32.79 | 43.2522 | 42.45 | 42.3333 | 39.4667 | 46.1167 | 42.4333 | 46.1904 | 33.7667 | 41.2667 | 72.52 | 41.6667 | 37.1167 | 45.2278 | 79.3833 | 98.75 | 35.72
| Median |  | 47.9355 | 52.95 | 58.2 | 85.704 | 42.3 | 32.795 | 43.3032 | 42.6 | 42.35 | 39.65 | 46.1 | 42.8 | 45.9683 | 33.9 | 41.315 | 72.87 | 41.5 | 36.9 | 45.033 | 76.75 | 98.6 | 35.9
| Std Dev. |  | 1.1945 | 2.3594 | 0.53914 | 0.025001 | 0.84178 | 0.86536 | 1.7826 | 1.6097 | 1.0289 | 0.73121 | 0.87958 | 1.1255 | 0.92756 | 0.40332 | 0.995 | 1.2412 | 0.8165 | 1.1566 | 0.51077 | 6.7565 | 5.35 | 1.1712
| Rel.Std.Dev. |  | 2.476% | 4.457% | 0.9285% | 0.02917% | 1.98% | 2.639% | 4.121% | 3.792% | 2.431% | 1.853% | 1.907% | 2.652% | 2.008% | 1.194% | 2.411% | 1.711% | 1.96% | 3.116% | 1.129% | 8.511% | 5.418% | 3.279%
| PDM3 |  | 12.54% | 23.49% | 35.47% | 99.94% | -0.8322% | -23.5% | 0.907% | -0.9644% | -1.237% | -7.925% | 7.59% | -1.003% | 7.762% | -21.22% | -3.725% | 69.19% | -2.792% | -13.41% | 5.516% | 85.2% | 130.4% | -16.67%
| Z-Score (Absolute) |  | 0.82 | 1.55 | 2.34 | 6.58 | 0.06 | 1.55 | 0.06 | 0.07 | 0.08 | 0.52 | 0.5 | 0.07 | 0.51 | 1.4 | 0.25 | 4.56 | 0.19 | 0.89 | 0.36 | 5.61 | 8.59 | 1.1
| Analytical results for Co in OREAS 553b (Certified Value 0.548 wt.%) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*MS | AR*AAS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 0.58685 | 0.56126 | 0.616 | > 0.5 | 0.50317 | 0.6425 | 0.5366 | 0.61271 | 0.57755 | 0.475 | 0.461 | 0.464 | 0.541 | 0.466 | 0.51611 | 0.5835 | 0.5426 | > 0.5 | 0.5907 | 0.539 | 0.54222 | 0.5904 | 0.55598 | 0.45 | 0.575 | 0.4026 | 1
| 1 | 2 | 0.57599 | 0.56275 | 0.61 | > 0.5 | 0.50165 | 0.6448 | 0.5335 | 0.6062 | 0.58066 | 0.474 | 0.467 | 0.466 | 0.525 | 0.464 | 0.54026 | 0.588 | 0.5609 | > 0.5 | 0.596 | 0.55 | 0.54283 | 0.5916 | 0.57593 | 0.496 | 0.534 | 0.4212 | 30
| 1 | 3 | 0.58044 | 0.57231 | 0.607 | > 0.5 | 0.50402 | 0.6593 | 0.5389 | 0.60542 | 0.57471 | 0.483 | 0.456 | 0.46 | 0.529 | 0.484 | 0.54157 | 0.593 | 0.5532 | > 0.5 | 0.5942 | 0.543 | 0.5238 | 0.5025 | 0.56734 | 0.517 | 0.563 | 0.4006 | 16
| 1 | 4 | 0.58204 | 0.56916 | 0.61 | > 0.5 | 0.50499 | 0.6801 | 0.5334 | 0.61191 | 0.57627 | 0.457 | 0.457 | 0.468 | 0.54 | 0.461 | 0.52563 | 0.595 | 0.5537 | > 0.5 | 0.6102 | 0.538 | 0.52819 | 0.5975 | 0.56883 | 0.511 | 0.566 | 0.3846 | 0.54777
| 1 | 5 | 0.58297 | 0.56392 | 0.614 | > 0.5 | 0.5013 | 0.6823 | 0.5379 | 0.60557 | 0.57313 | 0.465 | 0.452 | 0.482 | 0.539 | 0.461 | 0.51712 | 0.5925 | 0.5617 | > 0.5 | 0.5721 | 0.562 | 0.53045 | 0.5938 | 0.56429 | 0.482 | 0.565 | 0.3913 | 158
| 1 | 6 | 0.5898 | 0.54736 | 0.616 | > 0.5 | 0.50874 | 0.6873 | 0.5397 | 0.6062 | 0.58059 | 0.477 | 0.445 | 0.471 | 0.543 | 0.472 | 0.50424 | 0.5955 | 0.5636 | > 0.5 | 0.5904 | 0.545 | 0.52867 | 0.5953 | 0.56891 | 0.552 | 0.55 | 0.4205
| Mean |  | 0.58301 | 0.56279 | 0.61217 |  | 0.50398 | 0.66605 | 0.53667 | 0.608 | 0.57715 | 0.47183 | 0.45633 | 0.4685 | 0.53617 | 0.468 | 0.52415 | 0.59125 | 0.55595 |  | 0.59227 | 0.54617 | 0.53269 | 0.57852 | 0.56688 | 0.50133 | 0.55883 | 0.40347
| Median |  | 0.58251 | 0.56333 | 0.612 |  | 0.50359 | 0.6697 | 0.53725 | 0.6062 | 0.57691 | 0.4745 | 0.4565 | 0.467 | 0.5395 | 0.465 | 0.52138 | 0.59275 | 0.5573 |  | 0.59245 | 0.544 | 0.52956 | 0.5927 | 0.56809 | 0.5035 | 0.564 | 0.4016
| Std Dev. |  | 0.0048513 | 0.0086377 | 0.0037103 |  | 0.0027179 | 0.019823 | 0.0026986 | 0.0033619 | 0.003072 | 0.0093041 | 0.0075277 | 0.0075829 | 0.0073326 | 0.0088318 | 0.014668 | 0.0046341 | 0.007835 |  | 0.012263 | 0.0088863 | 0.0079259 | 0.037327 | 0.0065662 | 0.034477 | 0.01458 | 0.01495
| Rel.Std.Dev. |  | 0.8321% | 1.535% | 0.6061% |  | 0.5393% | 2.976% | 0.5029% | 0.553% | 0.5323% | 1.972% | 1.65% | 1.619% | 1.368% | 1.887% | 2.798% | 0.7838% | 1.409% |  | 2.071% | 1.627% | 1.488% | 6.452% | 1.158% | 6.877% | 2.609% | 3.705%
| PDM3 |  | 6.434% | 2.743% | 11.76% |  | -7.994% | 21.59% | -2.027% | 11% | 5.364% | -13.86% | -16.69% | -14.47% | -2.118% | -14.56% | -4.311% | 7.938% | 1.493% |  | 8.123% | -0.2926% | -2.752% | 5.613% | 3.489% | -8.477% | 2.02% | -26.34%
| Z-Score (Absolute) |  | 0.6 | 0.22 | 1.14 | NA | 0.88 | 2.15 | 0.27 | 1.06 | 0.49 | 1.48 | 1.77 | 1.54 | 0.28 | 1.55 | 0.5 | 0.75 | 0.09 | NA | 0.77 | 0.09 | 0.34 | 0.51 | 0.3 | 0.93 | 0.15 | 2.76
| Analytical results for Cr in OREAS 553b (Certified Value 33.9 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 20 | 21 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 2
| 1 | 1 | 37 | 32 | 38 | 40 | 40 | 36 | 30 | 32.3193 | 33 | 33 | 30 | 34 | 31 | 32.9019 | 30.2 | 34 | 35 | 35 | 32.349 | 37 | 54 | 32 |  |  |  |  | 1
| 1 | 2 | 39 | 33.8 | 40 | 41 | 41 | 36 | 29 | 34.6964 | 33 | 32 | 30 | 34 | 30 | 34.8804 | 29.3 | 34 | 32 | 35 | 32.65 | 82 | 46 | 32 |  |  |  |  | #N/A
| 1 | 3 | 38 | 33.7 | 37 | 41 | 39 | 36 | 31 | 33.6284 | 33 | 32 | 30 | 33 | 31 | 34.5638 | 31.5 | 35 | 33 | 35 | 32.134 | 40 | 50 | 32 |  |  |  |  | 16
| 1 | 4 | 37 | 32.4 | 43 | 41 | 41 | 36 | 29 | 35.384 | 32 | 32 | 30 | 34 | 30 | 33.4541 | 30.6 | 35 | 32 | 34 | 32.651 | 60 | 54 | 30 |  |  |  |  | 33.9115
| 1 | 5 | 39 | 31.2 | 35 | 40 | 40 | 36 | 31 | 31.9249 | 32 | 32 | 31 | 34 | 30 | 32.8304 | 30.2 | 35 | 33 | 34 | 32.48 | 40 | 54 | 29 |  |  |  |  | 159
| 1 | 6 | 38 | 31.9 | 37 | 41 | 41 | 36 | 30 | 33.3704 | 33 | 31 | 31 | 34 | 30 | 32.4499 | 29.5 | 35 | 33 | 35 | 32.415 | 76 | 53 | 32
| Mean |  | 38 | 32.5 | 38.3333 | 40.6667 | 40.3333 | 36 | 30 | 33.5539 | 32.6667 | 32 | 30.3333 | 33.8333 | 30.3333 | 33.5134 | 30.2167 | 34.6667 | 33 | 34.6667 | 32.4465 | 55.8333 | 51.8333 | 31.1667
| Median |  | 38 | 32.2 | 37.5 | 41 | 40.5 | 36 | 30 | 33.4994 | 33 | 32 | 30 | 34 | 30 | 33.178 | 30.2 | 35 | 33 | 35 | 32.4475 | 50 | 53.5 | 32
| Std Dev. |  | 0.89443 | 1.0431 | 2.8048 | 0.5164 | 0.8165 | 0 | 0.89443 | 1.332 | 0.5164 | 0.63246 | 0.5164 | 0.40825 | 0.5164 | 0.99469 | 0.79352 | 0.5164 | 1.0954 | 0.5164 | 0.19622 | 19.8234 | 3.2506 | 1.3292
| Rel.Std.Dev. |  | 2.354% | 3.209% | 7.317% | 1.27% | 2.024% | 0% | 2.981% | 3.97% | 1.581% | 1.976% | 1.702% | 1.207% | 1.702% | 2.968% | 2.626% | 1.49% | 3.32% | 1.49% | 0.6047% | 35.5% | 6.271% | 4.265%
| PDM3 |  | 12.06% | -4.162% | 13.04% | 19.92% | 18.94% | 6.159% | -11.53% | -1.055% | -3.671% | -5.637% | -10.55% | -0.2306% | -10.55% | -1.174% | -10.9% | 2.227% | -2.688% | 2.227% | -4.32% | 64.64% | 52.85% | -8.094%
| Z-Score (Absolute) |  | 1.25 | 0.29 | 1.34 | 1.99 | 1.9 | 0.69 | 0.99 | 0.01 | 0.24 | 0.43 | 0.89 | 0.08 | 0.89 | 0.01 | 0.93 | 0.32 | 0.15 | 0.32 | 0.3 | 6.22 | 5.11 | 0.66
| Analytical results for Cs in OREAS 553b (Certified Value 0.29 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 2
| 1 | 1 | 0.42 | 0.35 | 0.49 | 0.81031 | 0.35 | 0.25 | 0.2835 | 0.26 | 0.24 | 0.24 | 0.26 | 0.28 | 0.3789 | 0.32 | 0.24 | 0.2 | 0.42 | 0.14 | 0.8 | 0.35 | [6 empty] | 1
| 1 | 2 | 0.41 | 0.33 | 0.53 | 0.82771 | 0.39 | 0.28 | 0.32907 | 0.24 | 0.25 | 0.24 | 0.23 | 0.28 | 0.4022 | 0.32 | 0.26 | 0.2 | 0.446 | 0.14 | 0.7 | 0.36 | [6 empty] | 12
| 1 | 3 | 0.42 | 0.31 | 0.53 | 0.82238 | 0.37 | 0.25 | 0.31637 | 0.25 | 0.25 | 0.24 | 0.24 | 0.28 | 0.39508 | 0.33 | 0.25 | 0.2 | 0.428 | 0.14 | 0.9 | 0.36 | [6 empty] | 16
| 1 | 4 | 0.42 | 0.31 | 0.53 | 0.82548 | 0.37 | 0.28 | 0.3133 | 0.24 | 0.25 | 0.25 | 0.25 | 0.27 | 0.38366 | 0.29 | 0.26 | 0.2 | 0.456 | 0.14 | 0.9 | 0.35 | [6 empty] | 0.29441
| 1 | 5 | 0.42 | 0.26 | 0.51 | 0.80636 | 0.4 | 0.27 | 0.27791 | 0.24 | 0.25 | 0.22 | 0.24 | 0.26 | 0.37627 | 0.31 | 0.24 | 0.2 | 0.461 | 0.14 | 0.8 | 0.34 | [6 empty] | 160
| 1 | 6 | 0.42 | 0.33 | 0.52 | 0.80266 | 0.37 | 0.26 | 0.27262 | 0.24 | 0.23 | 0.25 | 0.24 | 0.27 | 0.39182 | 0.32 | 0.25 | 0.2 | 0.458 | 0.13 | 0.8 | 0.34
| Mean |  | 0.41833 | 0.315 | 0.51833 | 0.81581 | 0.375 | 0.265 | 0.29879 | 0.245 | 0.245 | 0.24 | 0.24333 | 0.27333 | 0.38799 | 0.315 | 0.25 | 0.2 | 0.44483 | 0.13833 | 0.81667 | 0.35
| Median |  | 0.42 | 0.32 | 0.525 | 0.81634 | 0.37 | 0.265 | 0.2984 | 0.24 | 0.25 | 0.24 | 0.24 | 0.275 | 0.38774 | 0.32 | 0.25 | 0.2 | 0.451 | 0.14 | 0.8 | 0.35
| Std Dev. |  | 0.0040825 | 0.030822 | 0.016021 | 0.010686 | 0.017607 | 0.013784 | 0.023627 | 0.0083666 | 0.0083666 | 0.010954 | 0.010328 | 0.008165 | 0.010055 | 0.013784 | 0.0089443 | 3.0405e-17 | 0.017093 | 0.0040825 | 0.075277 | 0.0089443
| Rel.Std.Dev. |  | 0.9759% | 9.785% | 3.091% | 1.31% | 4.695% | 5.202% | 7.908% | 3.415% | 3.415% | 4.564% | 4.244% | 2.987% | 2.592% | 4.376% | 3.578% | 1.52e-14% | 3.843% | 2.951% | 9.218% | 2.556%
| PDM3 |  | 42.09% | 6.994% | 76.06% | 177.1% | 27.37% | -9.989% | 1.489% | -16.78% | -16.78% | -18.48% | -17.35% | -7.159% | 31.79% | 6.994% | -15.08% | -32.07% | 51.09% | -53.01% | 177.4% | 18.88%
| Z-Score (Absolute) |  | 1.15 | 0.08 | 2.19 | 5.26 | 0.7 | 0.43 | 0.08 | 0.64 | 0.64 | 0.69 | 0.66 | 0.35 | 0.84 | 0.08 | 0.59 | 1.11 | 1.43 | 1.74 | 5.27 | 0.45
| Analytical results for Cu in OREAS 553b (Certified Value 3.6 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*MS | AR*AAS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*AAS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*AAS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g |  | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | > 2 | 3.704 | 3.61 | > 2 | 3.516 | > 1 | 3.6 | 3.63 | 3.5513 | 3.61 | 3.67 | 3.59 | 3.67 | 3.54 | 3.5968 | 3.563 | 3.54 | > 2 | 3.6552 | 3.495 | > 1 | 3.71 | 3.4537 | 3.51 | 3.7188 | 3.48 | 1
| 1 | 2 | > 2 | 3.677 | 3.63 | > 2 | 3.5205 | > 1 | 3.6 | 3.66 | 3.5409 | 3.62 | 3.68 | 3.56 | 3.71 | 3.51 | 3.6995 | 3.569 | 3.56 | > 2 | 3.6854 | 3.533 | > 1 | 3.74 | 3.5632 | 3.59 | 3.7058 | 3.45 | 31
| 1 | 3 | > 2 | 3.686 | 3.59 | > 2 | 3.5148 | > 1 | 3.62 | 3.61 | 3.545 | 3.58 | 3.68 | 3.63 | 3.66 | 3.48 | 3.6969 | 3.605 | 3.66 | > 2 | 3.687 | 3.504 | > 1 | 3.75 | 3.4888 | 3.63 | 3.692 | 3.48 | 16
| 1 | 4 | > 2 | 3.617 | 3.61 | > 2 | 3.5207 | > 1 | 3.63 | 3.66 | 3.5877 | 3.63 | 3.68 | 3.62 | 3.67 | 3.51 | 3.6121 | 3.608 | 3.64 | > 2 | 3.7548 | 3.488 | > 1 | 3.75 | 3.5608 | 3.61 | 3.6976 | 3.47 | 3.6039
| 1 | 5 | > 2 | 3.659 | 3.59 | > 2 | 3.5255 | > 1 | 3.59 | 3.62 | 3.5822 | 3.62 | 3.68 | 3.55 | 3.6 | 3.5 | 3.5657 | 3.579 | 3.59 | > 2 | 3.5715 | 3.516 | > 1 | 3.77 | 3.5105 | 3.5 | 3.7099 | 3.46 | 161
| 1 | 6 | > 2 | 3.712 | 3.63 | > 2 | 3.5181 | > 1 | 3.63 | 3.64 | 3.567 | 3.6 | 3.68 | 3.6 | 3.66 | 3.52 | 3.5282 | 3.595 | 3.69 | > 2 | 3.6555 | 3.519 | > 1 | 3.78 | 3.5235 | 3.63 | 3.7097 | 3.48
| Mean |  |  | 3.6758 | 3.61 |  | 3.5193 |  | 3.6117 | 3.6367 | 3.5624 | 3.61 | 3.6783 | 3.5917 | 3.6617 | 3.51 | 3.6165 | 3.5865 | 3.6133 |  | 3.6682 | 3.5092 |  | 3.75 | 3.5168 | 3.5783 | 3.7056 | 3.47
| Median |  |  | 3.6815 | 3.61 |  | 3.5193 |  | 3.61 | 3.635 | 3.5592 | 3.615 | 3.68 | 3.595 | 3.665 | 3.51 | 3.6045 | 3.587 | 3.615 |  | 3.6704 | 3.51 |  | 3.75 | 3.517 | 3.6 | 3.7078 | 3.475
| Std Dev. |  |  | 0.034499 | 0.017889 |  | 0.0038578 |  | 0.017224 | 0.020656 | 0.019709 | 0.017889 | 0.0040825 | 0.031885 | 0.035449 | 0.02 | 0.069457 | 0.018929 | 0.059217 |  | 0.059716 | 0.016654 |  | 0.024495 | 0.042283 | 0.058793 | 0.0095681 | 0.012649
| Rel.Std.Dev. |  |  | 0.9385% | 0.4955% |  | 0.1096% |  | 0.4769% | 0.568% | 0.5533% | 0.4955% | 0.111% | 0.8878% | 0.9681% | 0.5698% | 1.921% | 0.5278% | 1.639% |  | 1.628% | 0.4746% |  | 0.6532% | 1.202% | 1.643% | 0.2582% | 0.3645%
| PDM3 |  |  | 1.996% | 0.1692% |  | -2.348% |  | 0.2155% | 0.9091% | -1.153% | 0.1692% | 2.065% | -0.3395% | 1.603% | -2.606% | 0.3506% | -0.4829% | 0.2617% |  | 1.785% | -2.629% |  | 4.054% | -2.418% | -0.7095% | 2.822% | -3.715%
| Z-Score (Absolute) |  | NA | 0.86 | 0 | NA | 1.18 | NA | 0.02 | 0.35 | 0.62 | 0 | 0.89 | 0.24 | 0.67 | 1.31 | 0.09 | 0.31 | 0.04 | NA | 0.76 | 1.32 | NA | 1.83 | 1.22 | 0.41 | 1.25 | 1.83
| Analytical results for Dy in OREAS 553b (Certified Value 0.8 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 0.913 | 1 | 1.5386 | 0.87513 | 0.754 | 0.68 | 0.78 | [19 empty] | 1
| 1 | 2 | 0.893 | 1 | 1.519 | 0.93237 | 0.754 | 0.7 | 0.78 | [19 empty] | 13
| 1 | 3 | 0.868 | 0.9 | 1.5554 | 0.90373 | 0.754 | 0.7 | 0.78 | [19 empty] | 16
| 1 | 4 | 0.912 | 0.9 | 1.5548 | 0.87764 | 0.744 | 0.7 | 0.74 | [19 empty] | 0.79906
| 1 | 5 | 0.864 | 0.9 | 1.5306 | 0.89021 | 0.75 | 0.68 | 0.76 | [19 empty] | 162
| 1 | 6 | 0.894 | 0.9 | 1.5264 | 0.89872 | 0.744 | 0.69 | 0.76
| Mean |  | 0.89067 | 0.93333 | 1.5375 | 0.8963 | 0.75 | 0.69167 | 0.76667
| Median |  | 0.8935 | 0.9 | 1.5346 | 0.89447 | 0.752 | 0.695 | 0.77
| Std Dev. |  | 0.020954 | 0.05164 | 0.015036 | 0.020954 | 0.004899 | 0.0098319 | 0.01633
| Rel.Std.Dev. |  | 2.353% | 5.533% | 0.978% | 2.338% | 0.6532% | 1.421% | 2.13%
| PDM3 |  | 11.46% | 16.8% | 92.41% | 12.17% | -6.14% | -13.44% | -4.054%
| Z-Score (Absolute) |  | 0.57 | NA | 6.53 | 0.62 | 0.73 | 1.26 | 0.57
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 553b (Certified Value 0.36 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 3
| 1 | 1 | 0.391 | 0.5 | 0.80512 | 0.3838 | 0.331 | 0.32 | 0.37 | [19 empty] | 1
| 1 | 2 | 0.364 | 0.5 | 0.77048 | 0.4136 | 0.349 | 0.32 | 0.39 | [19 empty] | 14
| 1 | 3 | 0.37 | 0.4 | 0.78768 | 0.41404 | 0.339 | 0.33 | 0.38 | [19 empty] | 16
| 1 | 4 | 0.365 | 0.4 | 0.78136 | 0.37626 | 0.335 | 0.34 | 0.37 | [19 empty] | 0.36324
| 1 | 5 | 0.359 | 0.3 | 0.8016 | 0.38267 | 0.335 | 0.32 | 0.37 | [19 empty] | 163
| 1 | 6 | 0.382 | 0.4 | 0.78904 | 0.37323 | 0.343 | 0.33 | 0.38
| Mean |  | 0.37183 | 0.41667 | 0.78921 | 0.3906 | 0.33867 | 0.32667 | 0.37667
| Median |  | 0.3675 | 0.4 | 0.78836 | 0.38323 | 0.337 | 0.325 | 0.375
| Std Dev. |  | 0.012222 | 0.075277 | 0.012818 | 0.018411 | 0.0065013 | 0.008165 | 0.008165
| Rel.Std.Dev. |  | 3.287% | 18.07% | 1.624% | 4.713% | 1.92% | 2.499% | 2.168%
| PDM3 |  | 2.366% | 14.71% | 117.3% | 7.532% | -6.765% | -10.07% | 3.697%
| Z-Score (Absolute) |  | 0.09 | 0.71 | 7.32 | 0.25 | 0.67 | 0.89 | 0
| Analytical results for Eu in OREAS 553b (Certified Value 0.31 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 0.352 | 0.4 | 0.70784 | 0.3398 | 0.293 | 0.27 | 0.3 | [19 empty] | 1
| 1 | 2 | 0.345 | 0.4 | 0.67544 | 0.37398 | 0.297 | 0.27 | 0.3 | [19 empty] | 15
| 1 | 3 | 0.339 | 0.4 | 0.70208 | 0.37344 | 0.291 | 0.28 | 0.29 | [19 empty] | 16
| 1 | 4 | 0.355 | 0.4 | 0.69344 | 0.35431 | 0.298 | 0.28 | 0.29 | [19 empty] | 0.31443
| 1 | 5 | 0.351 | 0.4 | 0.70712 | 0.35716 | 0.306 | 0.26 | 0.28 | [19 empty] | 164
| 1 | 6 | 0.361 | 0.4 | 0.69224 | 0.35208 | 0.304 | 0.28 | 0.29
| Mean |  | 0.3505 | 0.4 | 0.69636 | 0.35846 | 0.29817 | 0.27333 | 0.29167
| Median |  | 0.3515 | 0.4 | 0.69776 | 0.35573 | 0.2975 | 0.275 | 0.29
| Std Dev. |  | 0.0076877 | 6.0809e-17 | 0.012194 | 0.013216 | 0.0059133 | 0.008165 | 0.0075277
| Rel.Std.Dev. |  | 2.193% | 1.52e-14% | 1.751% | 3.687% | 1.983% | 2.987% | 2.581%
| PDM3 |  | 11.47% | 27.22% | 121.5% | 14.01% | -5.171% | -13.07% | -7.238%
| Z-Score (Absolute) |  | 0.53 | NA | 7.51 | 0.69 | 0.53 | 1.03 | 0.66
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 553b (Certified Value 2.97 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 3.19 | 2.92 | 2.94 | 3.03 | 2.8793 | 3.31 | 2.96 | 3.08 | 2.9146 | 3.08 | 3.1 | 2.93 | 3.18 | 2.94 | 3.0056 | 3.06 | 2.93 | 3.07 | 2.9134 | 2.93 | 3 | 2.76 | 2.7154 | 3.455 | 2.94 | 2.92 | 1
| 1 | 2 | 3.31 | 2.93 | 3.09 | 3.03 | 2.8778 | 3.32 | 2.97 | 3.19 | 2.9013 | 3.04 | 3.12 | 2.94 | 3.11 | 2.96 | 3.0933 | 2.98 | 2.98 | 3.23 | 2.9061 | 2.83 | 2.99 | 2.79 | 2.8039 | 3.441 | 2.8 | 3 | #N/A
| 1 | 3 | 3.22 | 2.92 | 2.96 | 2.93 | 2.8765 | 3.48 | 2.97 | 3.12 | 3.0061 | 3.07 | 3.04 | 2.88 | 3.12 | 3 | 2.9982 | 3 | 2.97 | 3.15 | 2.9382 | 2.91 | 2.94 | 2.76 | 2.7462 | 3.539 | 2.69 | 2.89 | 16
| 1 | 4 | 3.23 | 2.87 | 3.11 | 3.01 | 2.8723 | 3.37 | 2.95 | 3.15 | 2.896 | 3 | 3.07 | 2.95 | 3.16 | 2.9 | 3.0324 | 3.01 | 2.99 | 3.2 | 2.9994 | 2.9 | 2.97 | 2.75 | 2.8047 | 3.476 | 2.87 | 2.89 | 2.9675
| 1 | 5 | 3.32 | 2.85 | 2.97 | 2.88 | 2.8746 | 3.48 | 2.98 | 3.09 | 3.0187 | 2.99 | 3.02 | 3.05 | 3.14 | 2.91 | 2.9736 | 3.05 | 2.93 | 3.22 | 2.8215 | 2.9 | 2.96 | 2.75 | 2.7536 | 3.497 | 3.04 | 2.93 | 165
| 1 | 6 | 3.28 | 2.85 | 3.04 | 3.02 | 2.8797 | 3.49 | 2.97 | 3.2 | 2.9626 | 3.06 | 3.04 | 2.99 | 3.15 | 2.89 | 2.9327 | 3.01 | 2.98 | 3.13 | 2.9363 | 2.92 | 2.96 | 2.79 | 2.7672 | 3.497 | 3.02 | 2.92
| Mean |  | 3.2583 | 2.89 | 3.0183 | 2.9833 | 2.8767 | 3.4083 | 2.9667 | 3.1383 | 2.9499 | 3.04 | 3.065 | 2.9567 | 3.1433 | 2.9333 | 3.0059 | 3.0183 | 2.9633 | 3.1667 | 2.9192 | 2.8983 | 2.97 | 2.7667 | 2.7652 | 3.4842 | 2.8933 | 2.925
| Median |  | 3.255 | 2.895 | 3.005 | 3.015 | 2.8772 | 3.425 | 2.97 | 3.135 | 2.9386 | 3.05 | 3.055 | 2.945 | 3.145 | 2.925 | 3.0019 | 3.01 | 2.975 | 3.175 | 2.9249 | 2.905 | 2.965 | 2.76 | 2.7604 | 3.4865 | 2.905 | 2.92
| Std Dev. |  | 0.052694 | 0.037417 | 0.071949 | 0.06314 | 0.0028841 | 0.084715 | 0.010328 | 0.050365 | 0.053956 | 0.037417 | 0.038859 | 0.05785 | 0.02582 | 0.041793 | 0.0544 | 0.030605 | 0.026583 | 0.061536 | 0.058028 | 0.035449 | 0.021909 | 0.018619 | 0.034752 | 0.034977 | 0.13441 | 0.040373
| Rel.Std.Dev. |  | 1.617% | 1.295% | 2.384% | 2.116% | 0.1003% | 2.486% | 0.3481% | 1.605% | 1.829% | 1.231% | 1.268% | 1.957% | 0.8214% | 1.425% | 1.81% | 1.014% | 0.8971% | 1.943% | 1.988% | 1.223% | 0.7377% | 0.673% | 1.257% | 1.004% | 4.646% | 1.38%
| PDM3 |  | 9.799% | -2.613% | 1.711% | 0.5321% | -3.061% | 14.85% | -0.02958% | 5.755% | -0.5953% | 2.442% | 3.284% | -0.3666% | 5.924% | -1.153% | 1.294% | 1.711% | -0.1419% | 6.71% | -1.631% | -2.332% | 0.08275% | -6.769% | -6.82% | 17.41% | -2.501% | -1.434%
| Z-Score (Absolute) |  | 2.76 | 0.75 | 0.48 | 0.14 | 0.87 | 4.19 | 0.02 | 1.62 | 0.18 | 0.68 | 0.92 | 0.11 | 1.67 | 0.33 | 0.36 | 0.48 | 0.05 | 1.89 | 0.47 | 0.67 | 0.02 | 1.92 | 1.93 | 4.91 | 0.71 | 0.41
| Analytical results for Ga in OREAS 553b (Certified Value 7.67 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 2
| 1 | 1 | 8.55 | 7.6 | 10.7 | 13.5 | 8.25 | 6.8 | 8.3369 | 7.52 | 6.77 | 7.09 | 8.14 | 7.24 | 8.1541 | 6.93 | 8.3 | 7 | 9 | 9 | 8.236 | 4.6 | 14.3 | 5.78 |  |  |  |  | 1
| 1 | 2 | 8.61 | 8.3 | 10.9 | 13.7 | 8.4 | 7.1 | 9.4766 | 7.5 | 7.02 | 7.3 | 7.96 | 7.17 | 8.3702 | 6.68 | 8.4 | 7 | 9 | 9 | 8.401 | 4.62 | 13.4 | 5.76 |  |  |  |  | 17
| 1 | 3 | 8.24 | 8.1 | 10.4 | 13.7 | 8.51 | 6.6 | 8.6017 | 7.43 | 6.98 | 6.78 | 8.1 | 7.36 | 8.4668 | 7.39 | 8.5 | 7 | 9 | 8.9 | 7.976 | 4.8 | 14.8 | 5.48 |  |  |  |  | 16
| 1 | 4 | 8.46 | 7.5 | 10.6 | 13.5 | 8.36 | 7.2 | 9.3638 | 6.73 | 6.9 | 7.09 | 8.34 | 7.3 | 8.2052 | 6.92 | 8.6 | 6 | 9 | 8.9 | 8.042 | 4.8 | 14.6 | 5.56 |  |  |  |  | 7.6749
| 1 | 5 | 8.32 | 6.7 | 9.93 | 13.3 | 8.11 | 6.8 | 8.5614 | 7.07 | 6.68 | 6.94 | 8.11 | 6.76 | 7.94 | 6.72 | 8.6 | 6 | 9 | 9.1 | 8.135 | 4.54 | 13.8 | 5.52 |  |  |  |  | 166
| 1 | 6 | 8.48 | 7.2 | 10.3 | 13.2 | 8.05 | 6.9 | 8.2346 | 7.21 | 6.58 | 6.77 | 8.6 | 6.9 | 7.7549 | 6.77 | 8.7 | 7 | 9 | 9 | 8.107 | 4.79 | 13.4 | 5.66
| Mean |  | 8.4433 | 7.5667 | 10.4717 | 13.4833 | 8.28 | 6.9 | 8.7625 | 7.2433 | 6.8217 | 6.995 | 8.2083 | 7.1217 | 8.1486 | 6.9017 | 8.5167 | 6.6667 | 9 | 8.9833 | 8.1495 | 4.6917 | 14.05 | 5.6267
| Median |  | 8.47 | 7.55 | 10.5 | 13.5 | 8.305 | 6.85 | 8.5815 | 7.32 | 6.835 | 7.015 | 8.125 | 7.205 | 8.1797 | 6.845 | 8.55 | 7 | 9 | 9 | 8.121 | 4.705 | 14.05 | 5.61
| Std Dev. |  | 0.13952 | 0.58538 | 0.34061 | 0.20412 | 0.17686 | 0.21909 | 0.52871 | 0.30722 | 0.17417 | 0.2054 | 0.22737 | 0.2387 | 0.26537 | 0.26042 | 0.1472 | 0.5164 | 0 | 0.075277 | 0.15127 | 0.11805 | 0.60581 | 0.12628
| Rel.Std.Dev. |  | 1.652% | 7.736% | 3.253% | 1.514% | 2.136% | 3.175% | 6.034% | 4.241% | 2.553% | 2.936% | 2.77% | 3.352% | 3.257% | 3.773% | 1.728% | 7.746% | 0% | 0.838% | 1.856% | 2.516% | 4.312% | 2.244%
| PDM3 |  | 10.01% | -1.411% | 36.44% | 75.68% | 7.884% | -10.1% | 14.17% | -5.624% | -11.12% | -8.859% | 6.95% | -7.209% | 6.171% | -10.08% | 10.97% | -13.14% | 17.26% | 17.05% | 6.183% | -38.87% | 83.06% | -26.69%
| Z-Score (Absolute) |  | 0.21 | 0.41 | 1.62 | 3.72 | 0.09 | 0.87 | 0.43 | 0.63 | 0.93 | 0.81 | 0.04 | 0.72 | 0 | 0.87 | 0.26 | 1.03 | 0.59 | 0.58 | 0 | 2.41 | 4.12 | 1.76
| Analytical results for Gd in OREAS 553b (Certified Value 1.74 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 1.725 | 2.1 | 4.3938 | 1.9436 | 1.635 | 1.38 | 1.74 | [19 empty] | 1
| 1 | 2 | 1.664 | 2.1 | 4.4394 | 2.0523 | 1.729 | 1.38 | 1.71 | [19 empty] | 18
| 1 | 3 | 1.694 | 2 | 4.4132 | 2.1909 | 1.613 | 1.44 | 1.68 | [19 empty] | 16
| 1 | 4 | 1.741 | 2 | 4.3902 | 1.9125 | 1.626 | 1.41 | 1.61 | [19 empty] | 1.7358
| 1 | 5 | 1.648 | 1.8 | 4.3619 | 1.9273 | 1.637 | 1.35 | 1.62 | [19 empty] | 167
| 1 | 6 | 1.723 | 2 | 4.3734 | 1.9902 | 1.667 | 1.38 | 1.67
| Mean |  | 1.6992 | 2 | 4.3953 | 2.0028 | 1.6512 | 1.39 | 1.6717
| Median |  | 1.7085 | 2 | 4.392 | 1.9669 | 1.636 | 1.38 | 1.675
| Std Dev. |  | 0.037059 | 0.10954 | 0.027904 | 0.10521 | 0.042097 | 0.030984 | 0.050365
| Rel.Std.Dev. |  | 2.181% | 5.477% | 0.6349% | 5.253% | 2.55% | 2.229% | 3.013%
| PDM3 |  | -2.11% | 15.22% | 153.2% | 15.38% | -4.876% | -19.92% | -3.695%
| Z-Score (Absolute) |  | 0 | 0.67 | 6.04 | 0.68 | 0.11 | 0.69 | 0.06
| Analytical results for Ge in OREAS 553b (Indicative Value 0.12 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [10 empty] | 2
| 1 | 1 | < 0.1 | 0.23 | 0.1 | 0.06 | 0.06 | 0.09 | 0.08 | 0.08 | < 0.2 | < 0.1 | 1.2 | 7.8 | 0.201 | < 0.01 | 0.14 | 0.09 | [10 empty] | 1
| 1 | 2 | < 0.1 | 0.25 | 0.2 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | < 0.2 | < 0.1 | 1.2 | 6.8 | 0.191 | < 0.01 | 0.14 | 0.09 | [10 empty] | 2
| 1 | 3 | < 0.1 | 0.24 | 0.1 | 0.05 | 0.07 | 0.07 | 0.08 | 0.07 | < 0.2 | < 0.1 | 1.2 | 7.8 | 0.207 | < 0.01 | 0.14 | 0.08 | [10 empty] | 16
| 1 | 4 | < 0.1 | 0.26 | 0.1 | 0.06 | 0.06 | 0.09 | 0.08 | 0.07 | < 0.2 | < 0.1 | 1.2 | 6.2 | 0.187 | < 0.01 | 0.13 | 0.08 | [10 empty] | 0.11772
| 1 | 5 | < 0.1 | 0.27 | 0.2 | 0.07 | 0.06 | 0.08 | 0.08 | 0.07 | < 0.2 | < 0.1 | 1.2 | 7.9 | 0.178 | < 0.01 | 0.14 | 0.09 | [10 empty] | 31
| 1 | 6 | < 0.1 | 0.26 | 0.2 | 0.06 | 0.05 | 0.08 | 0.08 | 0.07 | < 0.2 | < 0.1 | 1.2 | 7.1 | 0.209 | < 0.01 | 0.12 | 0.09
| Mean |  |  | 0.25167 | 0.15 | 0.061667 | 0.061667 | 0.081667 | 0.08 | 0.073333 |  |  | 1.2 | 7.2667 | 0.1955 |  | 0.135 | 0.086667
| Median |  |  | 0.255 | 0.15 | 0.06 | 0.06 | 0.08 | 0.08 | 0.07 |  |  | 1.2 | 7.45 | 0.196 |  | 0.14 | 0.09
| Std Dev. |  |  | 0.01472 | 0.054772 | 0.0075277 | 0.0075277 | 0.0075277 | 0 | 0.005164 |  |  | 0 | 0.68605 | 0.012194 |  | 0.0083666 | 0.005164
| Rel.Std.Dev. |  |  | 5.849% | 36.51% | 12.21% | 12.21% | 9.218% | 0% | 7.042% |  |  | 0% | 9.441% | 6.237% |  | 6.197% | 5.958%
| PDM3 |  |  | 113.8% | 27.42% | -47.61% | -47.61% | -30.62% | -32.04% | -37.7% |  |  | 919.4% | 6073% | 66.08% |  | 14.68% | -26.38%
| Z-Score (Absolute) |  | 0.67 | 3.31 | 1.3 | 0.44 | 0.44 | 0.05 | 0.08 | 0.21 | 0.31 | 0.67 | 22.02 | 141.75 | 2.2 | 1.56 | 1 | 0.05
| Analytical results for Hf in OREAS 553b (Certified Value 0.3 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 23 | 24 | 25 | 26 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [7 empty] | 3
| 1 | 1 | 0.31 | 0.3 | 0.3 | 0.26 | 0.35 | 0.32992 | 0.31 | 0.3 | 0.3 | 0.27 | 0.31 | 0.31552 | 0.3 | 0.29 | 0.3 | 0.314 | 0.28 | 0.47 | 0.31 | [7 empty] | 1
| 1 | 2 | 0.3 | 0.3 | 0.3 | 0.27 | 0.37 | 0.34649 | 0.29 | 0.31 | 0.3 | 0.26 | 0.31 | 0.34577 | 0.3 | 0.28 | 0.3 | 0.33 | 0.3 | 0.38 | 0.31 | [7 empty] | 22
| 1 | 3 | 0.31 | 0.2 | 0.29 | 0.27 | 0.32 | 0.34167 | 0.3 | 0.3 | 0.28 | 0.26 | 0.3 | 0.37594 | 0.36 | 0.29 | 0.3 | 0.33 | 0.29 | 0.49 | 0.3 | [7 empty] | 16
| 1 | 4 | 0.33 | 0.3 | 0.29 | 0.26 | 0.38 | 0.35827 | 0.29 | 0.31 | 0.28 | 0.26 | 0.29 | 0.34908 | 0.29 | 0.28 | 0.3 | 0.345 | 0.3 | 0.45 | 0.32 | [7 empty] | 0.30042
| 1 | 5 | 0.31 | 0.2 | 0.29 | 0.26 | 0.33 | 0.33998 | 0.28 | 0.31 | 0.28 | 0.26 | 0.29 | 0.32308 | 0.3 | 0.28 | 0.3 | 0.32 | 0.28 | 0.45 | 0.3 | [7 empty] | 168
| 1 | 6 | 0.32 | 0.2 | 0.31 | 0.24 | 0.36 | 0.35895 | 0.3 | 0.3 | 0.3 | 0.26 | 0.29 | 0.33158 | 0.31 | 0.29 | 0.3 | 0.318 | 0.29 | 0.42 | 0.3
| Mean |  | 0.31333 | 0.25 | 0.29667 | 0.26 | 0.35167 | 0.34588 | 0.295 | 0.305 | 0.29 | 0.26167 | 0.29833 | 0.34016 | 0.31 | 0.285 | 0.3 | 0.32617 | 0.29 | 0.44333 | 0.30667
| Median |  | 0.31 | 0.25 | 0.295 | 0.26 | 0.355 | 0.34408 | 0.295 | 0.305 | 0.29 | 0.26 | 0.295 | 0.33867 | 0.3 | 0.285 | 0.3 | 0.325 | 0.29 | 0.45 | 0.305
| Std Dev. |  | 0.010328 | 0.054772 | 0.008165 | 0.010954 | 0.023166 | 0.011244 | 0.010488 | 0.0054772 | 0.010954 | 0.0040825 | 0.0098319 | 0.021739 | 0.025298 | 0.0054772 | 0 | 0.011286 | 0.0089443 | 0.038816 | 0.008165
| Rel.Std.Dev. |  | 3.296% | 21.91% | 2.752% | 4.213% | 6.588% | 3.251% | 3.555% | 1.796% | 3.777% | 1.56% | 3.296% | 6.391% | 8.161% | 1.922% | 0% | 3.46% | 3.084% | 8.755% | 2.662%
| PDM3 |  | 4.298% | -16.78% | -1.249% | -13.45% | 17.06% | 15.13% | -1.804% | 1.525% | -3.468% | -12.9% | -0.6946% | 13.23% | 3.189% | -5.133% | -0.1398% | 8.57% | -3.468% | 47.57% | 2.079%
| Z-Score (Absolute) |  | 0.67 | 2.53 | 0.17 | 2.02 | 2.61 | 2.32 | 0.25 | 0.25 | 0.51 | 1.94 | 0.08 | 2.03 | 0.51 | 0.76 | 0 | 1.32 | 0.51 | 7.25 | 0.34
| Analytical results for Hg in OREAS 553b (Certified Value 0.05 ppm) | [27 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 3
| 1 | 1 | 0.039 | 0.02 | 0.03 | 0.07 | < 0.05 | 0.05 | 0.07 | 0.05 | 0.05 | 0.04 | < 1 | < 1 | 0.05 | 0.09 | 0.04 | 0.04 | < 5 | 0.06 | 0.06 | 0.06 | [6 empty] | 1
| 1 | 2 | 0.017 | 0.02 | 0.03 | 0.05 | < 0.05 | 0.06 | 0.06 | 0.05 | 0.05 | 0.07 | < 1 | < 1 | 0.05 | 0.07 | 0.05 | 0.052 | < 5 | 0.06 | 0.06 | 0.07 | [6 empty] | 23
| 1 | 3 | 0.037 | 0.04 | 0.04 | 0.06 | < 0.05 | 0.05 | 0.06 | 0.03 | 0.05 | 0.06 | < 1 | < 1 | 0.05 | 0.09 | 0.05 | 0.054 | < 5 | 0.06 | 0.06 | 0.07 | [6 empty] | 16
| 1 | 4 | 0.029 | 0.01 | 0.03 | 0.05 | < 0.05 | 0.04 | 0.05 | 0.05 | 0.04 | 0.05 | < 1 | < 1 | 0.05 | 0.08 | 0.05 | 0.05 | < 5 | 0.06 | 0.08 | 0.07 | [6 empty] | 0.050371
| 1 | 5 | 0.038 | 0.03 | 0.04 | 0.07 | < 0.05 | 0.05 | 0.06 | 0.04 | 0.05 | 0.05 | < 1 | < 1 | 0.05 | 0.12 | 0.05 | 0.043 | < 5 | 0.06 | 0.05 | 0.07 | [6 empty] | 169
| 1 | 6 | 0.044 | 0.04 | 0.03 | 0.06 | < 0.05 | 0.04 | 0.07 | 0.07 | 0.05 | 0.05 | < 1 | < 1 | 0.05 | 0.1 | 0.05 | 0.05 | < 5 | 0.05 | 0.06 | 0.07
| Mean |  | 0.034 | 0.026667 | 0.033333 | 0.06 |  | 0.048333 | 0.061667 | 0.048333 | 0.048333 | 0.053333 |  |  | 0.05 | 0.091667 | 0.048333 | 0.048167 |  | 0.058333 | 0.061667 | 0.068333
| Median |  | 0.0375 | 0.025 | 0.03 | 0.06 |  | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 |  |  | 0.05 | 0.09 | 0.05 | 0.05 |  | 0.06 | 0.06 | 0.07
| Std Dev. |  | 0.0096333 | 0.012111 | 0.005164 | 0.0089443 |  | 0.0075277 | 0.0075277 | 0.013292 | 0.0040825 | 0.010328 |  |  | 7.6012e-18 | 0.017224 | 0.0040825 | 0.0054559 |  | 0.0040825 | 0.0098319 | 0.0040825
| Rel.Std.Dev. |  | 28.33% | 45.41% | 15.49% | 14.91% |  | 15.57% | 12.21% | 27.5% | 8.447% | 19.36% |  |  | 1.52e-14% | 18.79% | 8.447% | 11.33% |  | 6.999% | 15.94% | 5.974%
| PDM3 |  | -32.5% | -47.06% | -33.82% | 19.12% |  | -4.046% | 22.42% | -4.046% | -4.046% | 5.881% |  |  | -0.7368% | 81.98% | -4.046% | -4.376% |  | 15.81% | 22.42% | 35.66%
| Z-Score (Absolute) |  | 1.19 | 1.69 | 1.24 | 0.56 | 1.8 | 0.22 | 0.67 | 0.22 | 0.22 | 0.11 | 30.23 | 30.23 | 0.11 | 2.7 | 0.22 | 0.24 | 165.09 | 0.45 | 0.67 | 1.12
| Analytical results for Ho in OREAS 553b (Certified Value 0.14 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 0.144 | 0.2 | 0.27104 | 0.13798 | 0.754 | 0.12 | 0.13 | [19 empty] | 1
| 1 | 2 | 0.142 | 0.2 | 0.27264 | 0.15541 | 0.754 | 0.12 | 0.14 | [19 empty] | 5
| 1 | 3 | 0.139 | 0.1 | 0.25704 | 0.17181 | 0.754 | 0.12 | 0.13 | [19 empty] | 16
| 1 | 4 | 0.147 | 0.1 | 0.26664 | 0.14868 | 0.744 | 0.12 | 0.13 | [19 empty] | 0.1392
| 1 | 5 | 0.143 | 0.2 | 0.27384 | 0.14978 | 0.75 | 0.12 | 0.13 | [19 empty] | 170
| 1 | 6 | 0.144 | 0.1 | 0.24632 | 0.14332 | 0.744 | 0.12 | 0.13
| Mean |  | 0.14317 | 0.15 | 0.26459 | 0.15116 | 0.75 | 0.12 | 0.13167
| Median |  | 0.1435 | 0.15 | 0.26884 | 0.14923 | 0.752 | 0.12 | 0.13
| Std Dev. |  | 0.0026394 | 0.054772 | 0.010835 | 0.011726 | 0.004899 | 0 | 0.0040825
| Rel.Std.Dev. |  | 1.844% | 36.51% | 4.095% | 7.757% | 0.6532% | 0% | 3.101%
| PDM3 |  | 2.85% | 7.759% | 90.08% | 8.594% | 438.8% | -13.79% | -5.411%
| Z-Score (Absolute) |  | 0.25 | 0 | 4.21 | 0.04 | 22.07 | 1.1 | 0.67
| Analytical results for In in OREAS 553b (Certified Value 1.54 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 25 | 26 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.1g | 0.4g | [7 empty] | 3
| 1 | 1 | 1.55 | 1.45 | 1.72 | 1.654 | 1.54 | 1.6443 | 1.585 | 1.665 | 1.535 | 1.455 | 1.66 | 1.5953 | 1.43 | 1.44 | 1.51 | 10.7 | 1.257 | 1.68 | 1.59 | [7 empty] | 1
| 1 | 2 | 1.55 | 1.53 | 1.73 | 1.625 | 1.46 | 1.6611 | 1.56 | 1.67 | 1.555 | 1.415 | 1.66 | 1.6098 | 1.45 | 1.49 | 1.49 | 10 | 1.317 | 1.56 | 1.624 | [7 empty] | 6
| 1 | 3 | 1.51 | 1.44 | 1.73 | 1.619 | 1.41 | 1.733 | 1.53 | 1.65 | 1.465 | 1.42 | 1.61 | 1.6691 | 1.45 | 1.49 | 1.51 | 10.9 | 1.254 | 1.7 | 1.56 | [7 empty] | 16
| 1 | 4 | 1.5 | 1.55 | 1.74 | 1.67 | 1.53 | 1.7145 | 1.45 | 1.67 | 1.52 | 1.46 | 1.64 | 1.5856 | 1.41 | 1.5 | 1.49 | 11.5 | 1.323 | 1.63 | 1.571 | [7 empty] | 1.5437
| 1 | 5 | 1.49 | 1.43 | 1.76 | 1.602 | 1.41 | 1.6353 | 1.5 | 1.6 | 1.515 | 1.475 | 1.55 | 1.5653 | 1.5 | 1.48 | 1.49 | 11.6 | 1.275 | 1.62 | 1.539 | [7 empty] | 171
| 1 | 6 | 1.56 | 1.46 | 1.74 | 1.615 | 1.52 | 1.6718 | 1.535 | 1.59 | 1.505 | 1.435 | 1.545 | 1.5766 | 1.48 | 1.51 | 1.5 | 10.2 | 1.301 | 1.63 | 1.49
| Mean |  | 1.5267 | 1.4767 | 1.7367 | 1.6308 | 1.4783 | 1.6767 | 1.5267 | 1.6408 | 1.5158 | 1.4433 | 1.6108 | 1.6003 | 1.4533 | 1.485 | 1.4983 | 10.8167 | 1.2878 | 1.6367 | 1.5623
| Median |  | 1.53 | 1.455 | 1.735 | 1.622 | 1.49 | 1.6664 | 1.5325 | 1.6575 | 1.5175 | 1.445 | 1.625 | 1.5904 | 1.45 | 1.49 | 1.495 | 10.8 | 1.288 | 1.63 | 1.5655
| Std Dev. |  | 0.030111 | 0.050465 | 0.013663 | 0.025795 | 0.059805 | 0.039068 | 0.047293 | 0.036389 | 0.0304 | 0.023805 | 0.052385 | 0.037002 | 0.03266 | 0.02429 | 0.0098319 | 0.65549 | 0.030069 | 0.049261 | 0.045663
| Rel.Std.Dev. |  | 1.972% | 3.417% | 0.7867% | 1.582% | 4.045% | 2.33% | 3.098% | 2.218% | 2.006% | 1.649% | 3.252% | 2.312% | 2.247% | 1.636% | 0.6562% | 6.06% | 2.335% | 3.01% | 2.923%
| PDM3 |  | -1.105% | -4.344% | 12.5% | 5.643% | -4.236% | 8.612% | -1.105% | 6.29% | -1.807% | -6.503% | 4.347% | 3.663% | -5.856% | -3.804% | -2.941% | 600.7% | -16.58% | 6.02% | 1.205%
| Z-Score (Absolute) |  | 0 | 0.46 | 1.92 | 0.95 | 0.44 | 1.37 | 0 | 1.05 | 0.1 | 0.76 | 0.77 | 0.67 | 0.67 | 0.38 | 0.26 | 85.11 | 2.19 | 1.01 | 0.33
| Analytical results for K in OREAS 553b (Certified Value 0.19 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.2139 | 0.18 | 0.26 | 0.31 | 0.199 | 0.2 | 0.19 | 0.20048 | 0.18 | 0.18 | 0.18 | 0.2 | 0.18 | 0.1942 | 0.16 | 0.19 | 0.4735 | 0.2 | 0.19 | 0.18 | 0.21312 | 0.2474 | 0.736 | 0.162 |  |  | 1
| 1 | 2 | 0.2213 | 0.2 | 0.28 | 0.305 | 0.205 | 0.2 | 0.18 | 0.22016 | 0.18 | 0.18 | 0.17 | 0.19 | 0.18 | 0.20252 | 0.16 | 0.19 | 0.4909 | 0.19 | 0.19 | 0.18 | 0.21786 | 0.2283 | 0.702 | 0.169 |  |  | #N/A
| 1 | 3 | 0.2149 | 0.2 | 0.27 | 0.291 | 0.199 | 0.2 | 0.19 | 0.20493 | 0.18 | 0.18 | 0.17 | 0.19 | 0.18 | 0.20514 | 0.16 | 0.19 | 0.4639 | 0.19 | 0.19 | 0.18 | 0.21219 | 0.2275 | 0.709 | 0.161 |  |  | 16
| 1 | 4 | 0.2142 | 0.19 | 0.24 | 0.308 | 0.201 | 0.19 | 0.19 | 0.22415 | 0.17 | 0.18 | 0.18 | 0.2 | 0.18 | 0.19832 | 0.16 | 0.19 | 0.494 | 0.18 | 0.18 | 0.18 | 0.21409 | 0.1992 | 0.788 | 0.156 |  |  | 0.19011
| 1 | 5 | 0.2219 | 0.17 | 0.24 | 0.301 | 0.199 | 0.19 | 0.18 | 0.19832 | 0.18 | 0.18 | 0.18 | 0.19 | 0.17 | 0.19369 | 0.17 | 0.19 | 0.5034 | 0.19 | 0.18 | 0.18 | 0.21556 | 0.2241 | 0.701 | 0.16 |  |  | 172
| 1 | 6 | 0.2219 | 0.18 | 0.26 | 0.303 | 0.203 | 0.19 | 0.18 | 0.20342 | 0.18 | 0.17 | 0.18 | 0.2 | 0.18 | 0.19082 | 0.16 | 0.19 | 0.4919 | 0.19 | 0.18 | 0.18 | 0.2169 | 0.2499 | 0.655 | 0.166
| Mean |  | 0.21802 | 0.18667 | 0.25833 | 0.303 | 0.201 | 0.195 | 0.185 | 0.20858 | 0.17833 | 0.17833 | 0.17667 | 0.195 | 0.17833 | 0.19745 | 0.16167 | 0.19 | 0.48627 | 0.19 | 0.185 | 0.18 | 0.21495 | 0.2294 | 0.71517 | 0.16233
| Median |  | 0.2181 | 0.185 | 0.26 | 0.304 | 0.2 | 0.195 | 0.185 | 0.20417 | 0.18 | 0.18 | 0.18 | 0.195 | 0.18 | 0.19626 | 0.16 | 0.19 | 0.4914 | 0.19 | 0.185 | 0.18 | 0.21483 | 0.2279 | 0.7055 | 0.1615
| Std Dev. |  | 0.0040538 | 0.012111 | 0.016021 | 0.0067231 | 0.0025298 | 0.0054772 | 0.0054772 | 0.010839 | 0.0040825 | 0.0040825 | 0.005164 | 0.0054772 | 0.0040825 | 0.005554 | 0.0040825 | 3.0405e-17 | 0.014627 | 0.0063246 | 0.0054772 | 3.0405e-17 | 0.0022071 | 0.01838 | 0.044206 | 0.0045898
| Rel.Std.Dev. |  | 1.859% | 6.488% | 6.202% | 2.219% | 1.259% | 2.809% | 2.961% | 5.197% | 2.289% | 2.289% | 2.923% | 2.809% | 2.289% | 2.813% | 2.525% | 1.6e-14% | 3.008% | 3.329% | 2.961% | 1.689e-14% | 1.027% | 8.012% | 6.181% | 2.827%
| PDM3 |  | 14.68% | -1.809% | 35.89% | 59.39% | 5.731% | 2.575% | -2.686% | 9.716% | -6.192% | -6.192% | -7.069% | 2.575% | -6.192% | 3.862% | -14.96% | -0.05548% | 155.8% | -0.05548% | -2.686% | -5.316% | 13.07% | 20.67% | 276.2% | -14.61%
| Z-Score (Absolute) |  | 1.21 | 0.28 | 3.13 | 5.26 | 0.4 | 0.12 | 0.36 | 0.77 | 0.67 | 0.67 | 0.75 | 0.12 | 0.67 | 0.24 | 1.47 | 0.12 | 13.98 | 0.12 | 0.36 | 0.59 | 1.07 | 1.76 | 24.88 | 1.44
| Analytical results for La in OREAS 553b (Certified Value 19 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 1
| 1 | 1 | 21.306 | 17.3 | 24.8 | 27.8 | 37.2632 | 19.5 | 18.9 | 14.7 | 20.9598 | 19.5 | 18.7 | 18.4 | 20.2 | 20 | 20.2513 | 15.8 | 18.1 | 31.2 | 19 | 18 | 13.8 | 18.04 | 43 | 42.9 | 16.7 |  | 1
| 1 | 2 | 21.015 | 18.8 | 24.8 | 27.3 | 37.1968 | 20.8 | 20 | 14.3 | 22.9227 | 19.9 | 19.3 | 19.2 | 18.9 | 19.8 | 21.3919 | 15.3 | 18.6 | 32.14 | 17 | 18 | 14.2 | 18.483 | 2 | 38.8 | 17 |  | 7
| 1 | 3 | 20.95 | 18.7 | 22.3 | 27.8 | 37.204 | 20 | 19.2 | 14.5 | 21.1269 | 19.2 | 19.5 | 17.9 | 19.4 | 19.2 | 21.1295 | 15.9 | 18.9 | 31.18 | 18 | 18 | 13.9 | 17.859 | 26 | 45.6 | 16.3 |  | 16
| 1 | 4 | 21.799 | 17.4 | 24.4 | 27.2 | 37.2296 | 20.4 | 18.9 | 14.3 | 22.7035 | 18.3 | 19 | 18.9 | 20.1 | 19.8 | 20.2651 | 15.7 | 19.2 | 31.98 | 17 | 18 | 14 | 17.836 | < 1 | 45 | 15.5 |  | 18.9644
| 1 | 5 | 21.204 | 16.3 | 21.3 | 27.1 | 37.2184 | 20.2 | 18.9 | 14.6 | 19.6315 | 18.4 | 18.6 | 18.4 | 19.5 | 18.6 | 20.2071 | 15.8 | 18.8 | 31.74 | 17 | 18 | 14.5 | 18.078 | < 1 | 41.3 | 15.9 |  | 173
| 1 | 6 | 22.015 | 17.3 | 22.9 | 27.8 | 37.264 | 20.4 | 18.8 | 14.4 | 20.2378 | 19.2 | 18.2 | 18.5 | 20 | 18.9 | 20.2503 | 15.6 | 19.1 | 32.39 | 18 | 18 | 13.9 | 18.247 | < 1 | 41.5 | 16.5
| Mean |  | 21.3815 | 17.6333 | 23.4167 | 27.5 | 37.2293 | 20.2167 | 19.1167 | 14.4667 | 21.2637 | 19.0833 | 18.8833 | 18.55 | 19.6833 | 19.3833 | 20.5826 | 15.6833 | 18.7833 | 31.7717 | 17.6667 | 18 | 14.05 | 18.0905 | 23.6667 | 42.5167 | 16.3167
| Median |  | 21.255 | 17.35 | 23.65 | 27.55 | 37.224 | 20.3 | 18.9 | 14.45 | 21.0434 | 19.2 | 18.85 | 18.45 | 19.75 | 19.5 | 20.2582 | 15.75 | 18.85 | 31.86 | 17.5 | 18 | 13.95 | 18.059 | 26 | 42.2 | 16.4
| Std Dev. |  | 0.43205 | 0.95429 | 1.4689 | 0.33466 | 0.028874 | 0.44008 | 0.4535 | 0.1633 | 1.3155 | 0.62423 | 0.47924 | 0.45056 | 0.50365 | 0.56716 | 0.53218 | 0.2137 | 0.39707 | 0.49785 | 0.8165 | 0 | 0.25884 | 0.24479 | 20.5994 | 2.5357 | 0.54559
| Rel.Std.Dev. |  | 2.021% | 5.412% | 6.273% | 1.217% | 0.07756% | 2.177% | 2.372% | 1.129% | 6.187% | 3.271% | 2.538% | 2.429% | 2.559% | 2.926% | 2.586% | 1.363% | 2.114% | 1.567% | 4.622% | 0% | 1.842% | 1.353% | 87.04% | 5.964% | 3.344%
| PDM3 |  | 12.75% | -7.019% | 23.48% | 45.01% | 96.31% | 6.603% | 0.8031% | -23.72% | 12.12% | 0.6273% | -0.4273% | -2.185% | 3.791% | 2.209% | 8.533% | -17.3% | -0.9546% | 67.53% | -6.843% | -5.085% | -25.91% | -4.608% | 24.8% | 124.2% | -13.96%
| Z-Score (Absolute) |  | 1.03 | 0.65 | 1.95 | 3.79 | 8.16 | 0.51 | 0.01 | 2.08 | 0.98 | 0 | 0.09 | 0.24 | 0.27 | 0.13 | 0.67 | 1.53 | 0.13 | 5.71 | 0.64 | 0.49 | 2.26 | 0.45 | 3.15 | 10.54 | 1.24
| Analytical results for Li in OREAS 553b (Certified Value 35.5 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 1
| 1 | 1 | 38.4 | 44.8 | 47.4 | 56.99 | 42 | 41 | 28 | 36.2878 | 31.6 | 30.7 | 28.9 | 36.5 | 30.5 | 36.4814 | 28.4 | 38 | 35 | 19.6 | 40.14 | 33 | 41.7 | 22.1 |  |  |  |  | 1
| 1 | 2 | 37.6 | 45.7 | 47.6 | 58.25 | 43 | 41 | 26 | 39.13 | 31.8 | 31.2 | 29.4 | 35 | 30.5 | 38.0097 | 28.1 | 38 | 36.7 | 20.1 | 39.94 | 33 | 39.1 | 23 |  |  |  |  | #N/A
| 1 | 3 | 37.4 | 37.8 | 47.3 | 56.97 | 43 | 42 | 29 | 36.5317 | 32.2 | 31 | 28.6 | 36.2 | 31.8 | 37.1752 | 32 | 38 | 34.5 | 19 | 38.56 | 33 | 41.5 | 21.1 |  |  |  |  | 16
| 1 | 4 | 38 | 46.6 | 47.8 | 56.09 | 43 | 41 | 27 | 39.5647 | 30 | 30.5 | 29.4 | 36.6 | 30.1 | 36.7101 | 28.4 | 39 | 36.1 | 20.5 | 39.61 | 36 | 40.2 | 20.6 |  |  |  |  | 35.4644
| 1 | 5 | 39.3 | 37.1 | 48 | 56.2 | 42 | 42 | 25 | 37.623 | 30.7 | 30.3 | 30.1 | 35.9 | 30.3 | 36.7581 | 28.8 | 38 | 33.4 | 20 | 39.36 | 31 | 40.9 | 21 |  |  |  |  | 174
| 1 | 6 | 38.7 | 37.9 | 47.4 | 58.12 | 43 | 41 | 28 | 36.8302 | 31.7 | 28.9 | 29.2 | 37 | 31 | 35.9656 | 27.8 | 38 | 36.7 | 19.1 | 39.25 | 36 | 41.8 | 22.5
| Mean |  | 38.2333 | 41.65 | 47.5833 | 57.1033 | 42.6667 | 41.3333 | 27.1667 | 37.6613 | 31.3333 | 30.4333 | 29.2667 | 36.2 | 30.7 | 36.85 | 28.9167 | 38.1667 | 35.4 | 19.7167 | 39.4767 | 33.6667 | 40.8667 | 21.7167
| Median |  | 38.2 | 41.35 | 47.5 | 56.98 | 43 | 41 | 27.5 | 37.2266 | 31.65 | 30.6 | 29.3 | 36.35 | 30.5 | 36.7341 | 28.4 | 38 | 35.55 | 19.8 | 39.485 | 33 | 41.2 | 21.6
| Std Dev. |  | 0.71181 | 4.4814 | 0.27142 | 0.91891 | 0.5164 | 0.5164 | 1.472 | 1.3881 | 0.81894 | 0.81894 | 0.51251 | 0.6957 | 0.61644 | 0.69234 | 1.5471 | 0.40825 | 1.3297 | 0.59133 | 0.56174 | 1.9664 | 1.052 | 0.95376
| Rel.Std.Dev. |  | 1.862% | 10.76% | 0.5704% | 1.609% | 1.21% | 1.249% | 5.418% | 3.686% | 2.614% | 2.691% | 1.751% | 1.922% | 2.008% | 1.879% | 5.35% | 1.07% | 3.756% | 2.999% | 1.423% | 5.841% | 2.574% | 4.392%
| PDM3 |  | 7.808% | 17.44% | 34.17% | 61.02% | 20.31% | 16.55% | -23.4% | 6.195% | -11.65% | -14.19% | -17.48% | 2.074% | -13.43% | 3.907% | -18.46% | 7.62% | -0.1816% | -44.4% | 11.31% | -5.069% | 15.23% | -38.76%
| Z-Score (Absolute) |  | 0.22 | 0.67 | 1.45 | 2.69 | 0.8 | 0.63 | 1.22 | 0.15 | 0.68 | 0.8 | 0.95 | 0.04 | 0.76 | 0.04 | 0.99 | 0.21 | 0.15 | 2.2 | 0.39 | 0.37 | 0.57 | 1.94
| Analytical results for Lu in OREAS 553b (Certified Value 0.041 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 03 | 05 | 08 | 15 | 17 | 22 | 23 | 26 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | [17 empty] | 3
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.4g | [17 empty] | 3
| 1 | 1 | 0.045 | < 0.1 | 0.08112 | 0.04 | 0.052042 | 0.04 | < 5 | 0.038 | 0.04 | [17 empty] | 1
| 1 | 2 | 0.044 | < 0.1 | 0.08472 | 0.04 | 0.064546 | 0.04 | < 5 | 0.04 | 0.04 | [17 empty] | 8
| 1 | 3 | 0.045 | < 0.1 | 0.09032 | 0.04 | 0.059135 | 0.04 | < 5 | 0.04 | 0.04 | [17 empty] | 16
| 1 | 4 | 0.048 | < 0.1 | 0.08672 | 0.04 | 0.050272 | 0.04 | < 5 | 0.04 | 0.04 | [17 empty] | 0.041033
| 1 | 5 | 0.045 | < 0.1 | 0.08704 | 0.04 | 0.057346 | 0.04 | < 5 | 0.038 | 0.04 | [17 empty] | 175
| 1 | 6 | 0.047 | < 0.1 | 0.07736 | 0.04 | < 0.05 | 0.04 | < 5 | 0.041 | 0.04
| Mean |  | 0.045667 |  | 0.084547 | 0.04 | 0.056668 | 0.04 |  | 0.0395 | 0.04
| Median |  | 0.045 |  | 0.08572 | 0.04 | 0.057346 | 0.04 |  | 0.04 | 0.04
| Std Dev. |  | 0.0015055 |  | 0.0046409 | 0 | 0.0057207 | 0 |  | 0.0012247 | 0
| Rel.Std.Dev. |  | 3.297% |  | 5.489% | 0% | 10.1% | 0% |  | 3.101% | 0%
| PDM3 |  | 11.29% |  | 106% | -2.518% | 38.1% | -2.518% |  | -3.737% | -2.518%
| Z-Score (Absolute) |  | 0 | 0.52 | 4.63 | 0.67 | 0.68 | 0.67 | 292.06 | 0.73 | 0.67
| Analytical results for Mg in OREAS 553b (Certified Value 2.85 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 3.01 | 2.76 | 2.97 | 3.02 | 2.996 | 2.89 | 2.54 | 2.7739 | 2.78 | 2.8 | 2.65 | 3.07 | 2.66 | 3.0078 | 2.78 | 2.92 | 3.1 | 2.85 | 2.92 | 2.35 | 2.7292 | 2.846 | 3.21 | 2.655 |  |  | 1
| 1 | 2 | 3.11 | 2.88 | 2.97 | 3.09 | 3.019 | 2.94 | 2.51 | 2.8519 | 2.76 | 2.82 | 2.65 | 3.02 | 2.64 | 3.0296 | 2.73 | 2.94 | 3.27 | 2.7 | 2.88 | 2.37 | 2.8105 | 2.865 | 2.96 | 2.756 |  |  | #N/A
| 1 | 3 | 3.03 | 2.88 | 3.14 | 3.04 | 3.065 | 2.9 | 2.53 | 2.8398 | 2.78 | 2.77 | 2.62 | 3.01 | 2.7 | 2.9512 | 2.76 | 2.94 | 3.22 | 2.75 | 2.86 | 2.37 | 2.7326 | 2.913 | 2.92 | 2.619 |  |  | 16
| 1 | 4 | 3.01 | 2.77 | 3.08 | 3.05 | 3.041 | 2.88 | 2.55 | 2.8835 | 2.65 | 2.78 | 2.66 | 3.06 | 2.6 | 2.9475 | 2.74 | 2.99 | 3.27 | 2.68 | 2.83 | 2.36 | 2.7732 | 2.834 | 3.11 | 2.534 |  |  | 2.8486
| 1 | 5 | 3.11 | 2.65 | 3 | 2.99 | 3.052 | 2.91 | 2.5 | 2.7634 | 2.68 | 2.75 | 2.74 | 3.02 | 2.59 | 3.0206 | 2.81 | 2.94 | 3.32 | 2.69 | 2.82 | 2.37 | 2.7563 | 2.84 | 3.28 | 2.582 |  |  | 176
| 1 | 6 | 3.09 | 2.71 | 3.12 | 3.02 | 3.112 | 2.9 | 2.53 | 2.8437 | 2.76 | 2.67 | 2.71 | 3.07 | 2.63 | 2.946 | 2.75 | 2.98 | 3.29 | 2.76 | 2.83 | 2.35 | 2.7609 | 2.973 | 3.22 | 2.701
| Mean |  | 3.06 | 2.775 | 3.0467 | 3.035 | 3.0475 | 2.9033 | 2.5267 | 2.826 | 2.735 | 2.765 | 2.6717 | 3.0417 | 2.6367 | 2.9838 | 2.7617 | 2.9517 | 3.245 | 2.7383 | 2.8567 | 2.3617 | 2.7605 | 2.8785 | 3.1167 | 2.6412
| Median |  | 3.06 | 2.765 | 3.04 | 3.03 | 3.0465 | 2.9 | 2.53 | 2.8417 | 2.76 | 2.775 | 2.655 | 3.04 | 2.635 | 2.9795 | 2.755 | 2.94 | 3.27 | 2.725 | 2.845 | 2.365 | 2.7586 | 2.8555 | 3.16 | 2.637
| Std Dev. |  | 0.04858 | 0.091815 | 0.076333 | 0.033912 | 0.039984 | 0.020656 | 0.018619 | 0.047156 | 0.055767 | 0.05244 | 0.04446 | 0.027869 | 0.040332 | 0.039593 | 0.029269 | 0.027142 | 0.078166 | 0.063692 | 0.038297 | 0.0098319 | 0.029756 | 0.054445 | 0.14787 | 0.080542
| Rel.Std.Dev. |  | 1.588% | 3.309% | 2.505% | 1.117% | 1.312% | 0.7115% | 0.7369% | 1.669% | 2.039% | 1.897% | 1.664% | 0.9162% | 1.53% | 1.327% | 1.06% | 0.9195% | 2.409% | 2.326% | 1.341% | 0.4163% | 1.078% | 1.891% | 4.745% | 3.049%
| PDM3 |  | 7.422% | -2.583% | 6.954% | 6.544% | 6.983% | 1.922% | -11.3% | -0.7914% | -3.987% | -2.934% | -6.211% | 6.778% | -7.439% | 4.747% | -3.051% | 3.619% | 13.92% | -3.87% | 0.2839% | -17.09% | -3.093% | 1.05% | 9.411% | -7.281%
| Z-Score (Absolute) |  | 0.94 | 0.29 | 0.89 | 0.84 | 0.89 | 0.27 | 1.36 | 0.07 | 0.46 | 0.33 | 0.73 | 0.87 | 0.88 | 0.62 | 0.34 | 0.48 | 1.74 | 0.45 | 0.07 | 2.07 | 0.35 | 0.16 | 1.19 | 0.86
| Analytical results for Mn in OREAS 553b (Certified Value 0.013 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.0134 | 0.0129 | 0.0122 | 0.012 | 0.0149 | 0.0122 | 0.012 | 0.012899 | 0.0125 | 0.0133 | 0.0121 | 0.0129 | 0.0124 | 0.012743 | 0.0118 | 0.0124 | 0.0135 | 0.0133 | 0.0119 | 0.01177 | 0.011013 | 0.0139 | 0.0117 | 0.0127 |  |  | 1
| 1 | 2 | 0.0138 | 0.0132 | 0.0126 | 0.012 | 0.0148 | 0.0126 | 0.0129 | 0.013001 | 0.0124 | 0.0133 | 0.012 | 0.0127 | 0.0123 | 0.013142 | 0.0117 | 0.0126 | 0.0142 | 0.0128 | 0.0118 | 0.01202 | 0.011265 | 0.0139 | 0.0106 | 0.0129 |  |  | #N/A
| 1 | 3 | 0.0135 | 0.0131 | 0.0114 | 0.0115 | 0.0151 | 0.0125 | 0.0118 | 0.013075 | 0.0125 | 0.013 | 0.012 | 0.0126 | 0.0124 | 0.01295 | 0.0124 | 0.0126 | 0.0139 | 0.0132 | 0.0116 | 0.01206 | 0.011032 | 0.0139 | 0.0107 | 0.0125 |  |  | 16
| 1 | 4 | 0.0136 | 0.0128 | 0.0128 | 0.0122 | 0.015 | 0.0124 | 0.0121 | 0.013035 | 0.0121 | 0.0132 | 0.0121 | 0.0128 | 0.0121 | 0.01295 | 0.012 | 0.0126 | 0.0141 | 0.0132 | 0.0114 | 0.01186 | 0.011288 | 0.0132 | 0.0123 | 0.0125 |  |  | 0.012506
| 1 | 5 | 0.0138 | 0.0126 | 0.0115 | 0.0116 | 0.0152 | 0.012 | 0.0127 | 0.012896 | 0.0121 | 0.013 | 0.0126 | 0.0127 | 0.012 | 0.01253 | 0.012 | 0.0124 | 0.0141 | 0.013 | 0.0115 | 0.01196 | 0.011163 | 0.0139 | 0.0117 | 0.0126 |  |  | 177
| 1 | 6 | 0.0138 | 0.0125 | 0.0128 | 0.0119 | 0.0154 | 0.0121 | 0.0119 | 0.013144 | 0.0124 | 0.0129 | 0.0125 | 0.0128 | 0.0121 | 0.012463 | 0.0115 | 0.0126 | 0.0139 | 0.0133 | 0.0117 | 0.01171 | 0.011094 | 0.0139 | 0.0116 | 0.0126
| Mean |  | 0.01365 | 0.01285 | 0.012217 | 0.011867 | 0.015067 | 0.0123 | 0.012233 | 0.013008 | 0.012333 | 0.013117 | 0.012217 | 0.01275 | 0.012217 | 0.012796 | 0.0119 | 0.012533 | 0.01395 | 0.013133 | 0.01165 | 0.011897 | 0.011142 | 0.013783 | 0.011433 | 0.012633
| Median |  | 0.0137 | 0.01285 | 0.0124 | 0.01195 | 0.01505 | 0.0123 | 0.01205 | 0.013018 | 0.0124 | 0.0131 | 0.0121 | 0.01275 | 0.0122 | 0.012846 | 0.0119 | 0.0126 | 0.014 | 0.0132 | 0.01165 | 0.01191 | 0.011128 | 0.0139 | 0.01165 | 0.0126
| Std Dev. |  | 0.00017607 | 0.00027386 | 0.00063377 | 0.00026583 | 0.00021602 | 0.00023664 | 0.00045461 | 9.8302e-05 | 0.00018619 | 0.00017224 | 0.00026394 | 0.00010488 | 0.00017224 | 0.00026517 | 0.00030984 | 0.00010328 | 0.000251 | 0.00019664 | 0.00018708 | 0.0001401 | 0.00011673 | 0.00028577 | 0.00065625 | 0.00015055
| Rel.Std.Dev. |  | 1.29% | 2.131% | 5.188% | 2.24% | 1.434% | 1.924% | 3.716% | 0.7557% | 1.51% | 1.313% | 2.161% | 0.8226% | 1.41% | 2.072% | 2.604% | 0.824% | 1.799% | 1.497% | 1.606% | 1.178% | 1.048% | 2.073% | 5.74% | 1.192%
| PDM3 |  | 9.148% | 2.751% | -2.314% | -5.112% | 20.48% | -1.647% | -2.18% | 4.015% | -1.381% | 4.883% | -2.314% | 1.951% | -2.314% | 2.321% | -4.846% | 0.2186% | 11.55% | 5.016% | -6.845% | -4.872% | -10.9% | 10.21% | -8.577% | 1.018%
| Z-Score (Absolute) |  | 1.53 | 0.53 | 0.27 | 0.71 | 3.32 | 0.17 | 0.25 | 0.72 | 0.13 | 0.86 | 0.27 | 0.4 | 0.27 | 0.46 | 0.67 | 0.13 | 1.91 | 0.88 | 0.99 | 0.68 | 1.63 | 1.74 | 1.26 | 0.25
| Analytical results for Mo in OREAS 553b (Certified Value 14.2 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 2
| 1 | 1 | 15.1 | 14.21 | 14.3 | 14 | 14 | 15.1 | 12.9719 | 15.6 | 14.1 | 15.1 | 14.1 | 14.5 | 14.3087 | 13.87 | 13.71 | 14.24 | 14.1 | 13.9 | 13.9 | 6.9 | 13.728 | 11.8 | 14.1 | 16 |  |  | 1
| 1 | 2 | 15 | 14.31 | 14.7 | 14 | 14 | 15.26 | 13.0739 | 15.6 | 14.15 | 15.2 | 13.55 | 14.45 | 14.8707 | 13.01 | 14.15 | 14.86 | 14.1 | 13 | 14.1 | 7 | 14.128 | 12 | 14.1 | 16 |  |  | 9
| 1 | 3 | 15.2 | 13.7 | 14.2 | 14.6 | 14 | 14.83 | 13.3042 | 14.8 | 14.1 | 14.6 | 13.65 | 14.25 | 14.7636 | 13.56 | 14.15 | 14.33 | 14.2 | 13.5 | 14 | 7.2 | 13.623 | 12.1 | 14.6 | 17 |  |  | 16
| 1 | 4 | 15.4 | 14.21 | 14.9 | 14.3 | 14 | 14.8 | 13.3068 | 14.55 | 14.1 | 14.95 | 14.25 | 14.25 | 14.0884 | 13.32 | 14.16 | 14.74 | 14.4 | 13.5 | 13.8 | 6.9 | 14.128 | 12.2 | 13.7 | 16 |  |  | 14.1916
| 1 | 5 | 15.1 | 13.54 | 14.3 | 14.2 | 14 | 15.27 | 12.8632 | 14.8 | 13.45 | 15 | 14.05 | 13.15 | 14.3472 | 13.15 | 14.06 | 14.83 | 13.7 | 13.2 | 13.9 | 7.2 | 13.956 | 11.7 | 14.2 | 16 |  |  | 178
| 1 | 6 | 15.2 | 13.52 | 14.6 | 14.5 | 14 | 15.02 | 13.4006 | 14.95 | 13.45 | 15.1 | 14.25 | 13.3 | 14.2858 | 13.24 | 14.19 | 14.79 | 14.1 | 13.5 | 14 | 7.2 | 13.986 | 12 | 13.9 | 16
| Mean |  | 15.1667 | 13.915 | 14.5 | 14.2667 | 14 | 15.0467 | 13.1534 | 15.05 | 13.8917 | 14.9917 | 13.975 | 13.9833 | 14.4441 | 13.3583 | 14.07 | 14.6317 | 14.1 | 13.4333 | 13.95 | 7.0667 | 13.9248 | 11.9667 | 14.1 | 16.1667
| Median |  | 15.15 | 13.955 | 14.45 | 14.25 | 14 | 15.06 | 13.189 | 14.875 | 14.1 | 15.05 | 14.075 | 14.25 | 14.328 | 13.28 | 14.15 | 14.765 | 14.1 | 13.5 | 13.95 | 7.1 | 13.971 | 12 | 14.1 | 16
| Std Dev. |  | 0.13663 | 0.36687 | 0.27568 | 0.25033 | 0 | 0.20334 | 0.21487 | 0.44497 | 0.34266 | 0.21075 | 0.3029 | 0.59805 | 0.30443 | 0.3107 | 0.18166 | 0.27301 | 0.22804 | 0.30768 | 0.10488 | 0.15055 | 0.20838 | 0.18619 | 0.30332 | 0.40825
| Rel.Std.Dev. |  | 0.9008% | 2.636% | 1.901% | 1.755% | 0% | 1.351% | 1.634% | 2.957% | 2.467% | 1.406% | 2.167% | 4.277% | 2.108% | 2.326% | 1.291% | 1.866% | 1.617% | 2.29% | 0.7518% | 2.13% | 1.496% | 1.556% | 2.151% | 2.525%
| PDM3 |  | 6.87% | -1.949% | 2.173% | 0.5287% | -1.35% | 6.025% | -7.316% | 6.048% | -2.114% | 5.637% | -1.526% | -1.468% | 1.779% | -5.872% | -0.8571% | 3.101% | -0.6457% | -5.343% | -1.703% | -50.21% | -1.88% | -15.68% | -0.6457% | 13.92%
| Z-Score (Absolute) |  | 1.75 | 0.19 | 0.72 | 0.36 | 0.05 | 1.56 | 1.36 | 1.57 | 0.22 | 1.48 | 0.09 | 0.08 | 0.63 | 1.04 | 0.05 | 0.92 | 0.1 | 0.93 | 0.13 | 10.75 | 0.17 | 3.19 | 0.1 | 3.29
| Analytical results for Na in OREAS 553b (Certified Value 0.01 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | [6 empty] | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | [6 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 3
| 1 | 1 | 0.003 | 0.016 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.006 | 0.02 | < 0.01 | 0.0073067 | < 0.007 | 0.0065 | 0.009 | [6 empty] | 1
| 1 | 2 | 0.003 | 0.017 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.006 | 0.01 | < 0.01 | 0.0077961 | < 0.007 | 0.0072 | 0.009 | [6 empty] | 20
| 1 | 3 | 0.003 | 0.011 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.006 | 0.01 | < 0.01 | 0.007727 | < 0.007 | 0.0064 | 0.011 | [6 empty] | 16
| 1 | 4 | 0.003 | 0.016 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.01 | 0.0077943 | < 0.007 | 0.0064 | 0.009 | [6 empty] | 0.0097809
| 1 | 5 | 0.003 | 0.014 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.006 | 0.01 | < 0.01 | 0.0085769 | < 0.007 | 0.0065 | 0.011 | [6 empty] | 179
| 1 | 6 | 0.002 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.006 | 0.01 | < 0.01 | 0.0086516 | < 0.007 | 0.0067 | 0.01
| Mean |  | 0.0028333 | 0.014 | 0.01 | 0.01 | 0.01 |  |  |  | 0.011667 | 0.01 | 0.01 |  | 0.01 | 0.006 | 0.012 |  | 0.0079754 |  | 0.0066167 | 0.0098333
| Median |  | 0.003 | 0.015 | 0.01 | 0.01 | 0.01 |  |  |  | 0.01 | 0.01 | 0.01 |  | 0.01 | 0.006 | 0.01 |  | 0.0077952 |  | 0.0065 | 0.0095
| Std Dev. |  | 0.00040825 | 0.0028983 | 0 | 0 | 0 |  |  |  | 0.0040825 | 0 | 0 |  | 0 | 9.5015e-19 | 0.0044721 |  | 0.0005278 |  | 0.00030605 | 0.00098319
| Rel.Std.Dev. |  | 14.41% | 20.7% | 0% | 0% | 0% |  |  |  | 34.99% | 0% | 0% |  | 0% | 1.584e-14% | 37.27% |  | 6.618% |  | 4.625% | 9.999%
| PDM3 |  | -71.03% | 43.14% | 2.24% | 2.24% | 2.24% |  |  |  | 19.28% | 2.24% | 2.24% |  | 2.24% | -38.66% | 22.69% |  | -18.46% |  | -32.35% | 0.5363%
| Z-Score (Absolute) |  | 1.17 | 1.74 | 0.17 | 0.7 | 0.7 | 0.6 | 0.6 | 0.6 | 1.13 | 0.7 | 0.7 | 0.6 | 0.7 | 0.34 | 0.91 | 0.6 | 0.17 | 0.94 | 0.18 | 0.65
| Analytical results for Nb in OREAS 553b (Indicative Value 0.086 ppm) | [27 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 21 | 22 | 24 | 25 | 26 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [8 empty] | 3
| 1 | 1 | 0.07 | < 0.1 | < 0.2 | 0.24 | 0.13 | < 0.2 | 0.07 | 0.08 | 0.1 | 0.06 | 0.08 | 0.08959 | < 0.05 | < 0.1 | < 5 | 0.07 | 0.13 | 0.12 | [8 empty] | 1
| 1 | 2 | 0.07 | < 0.1 | < 0.2 | 0.24 | 0.11 | < 0.2 | 0.06 | 0.08 | 0.08 | 0.05 | 0.08 | 0.11192 | < 0.05 | < 0.1 | < 5 | 0.06 | 0.12 | 0.12 | [8 empty] | 26
| 1 | 3 | 0.07 | < 0.1 | < 0.2 | 0.23 | 0.13 | < 0.2 | 0.06 | 0.08 | 0.09 | 0.05 | 0.07 | 0.099814 | < 0.05 | < 0.1 | < 5 | 0.06 | 0.12 | 0.11 | [8 empty] | 16
| 1 | 4 | 0.08 | < 0.1 | < 0.2 | 0.24 | 0.12 | < 0.2 | 0.06 | 0.08 | 0.09 | 0.06 | 0.08 | 0.097499 | < 0.05 | < 0.1 | < 5 | 0.06 | 0.11 | 0.12 | [8 empty] | 0.086267
| 1 | 5 | 0.07 | < 0.1 | < 0.2 | 0.24 | 0.11 | < 0.2 | 0.06 | 0.08 | 0.08 | 0.06 | 0.07 | 0.09182 | < 0.05 | < 0.1 | < 5 | 0.06 | 0.12 | 0.12 | [8 empty] | 32
| 1 | 6 | 0.08 | < 0.1 | < 0.2 | 0.28 | 0.12 | < 0.2 | 0.06 | 0.07 | 0.1 | 0.06 | 0.08 | 0.092957 | < 0.05 | < 0.1 | < 5 | 0.05 | 0.11 | 0.11
| Mean |  | 0.073333 |  |  | 0.245 | 0.12 |  | 0.061667 | 0.078333 | 0.09 | 0.056667 | 0.076667 | 0.097267 |  |  |  | 0.06 | 0.11833 | 0.11667
| Median |  | 0.07 |  |  | 0.24 | 0.12 |  | 0.06 | 0.08 | 0.09 | 0.06 | 0.08 | 0.095228 |  |  |  | 0.06 | 0.12 | 0.12
| Std Dev. |  | 0.005164 |  |  | 0.017607 | 0.0089443 |  | 0.0040825 | 0.0040825 | 0.0089443 | 0.005164 | 0.005164 | 0.0081058 |  |  |  | 0.0063246 | 0.0075277 | 0.005164
| Rel.Std.Dev. |  | 7.042% |  |  | 7.186% | 7.454% |  | 6.62% | 5.212% | 9.938% | 9.113% | 6.736% | 8.334% |  |  |  | 10.54% | 6.361% | 4.426%
| PDM3 |  | -14.99% |  |  | 184% | 39.1% |  | -28.52% | -9.196% | 4.328% | -34.31% | -11.13% | 12.75% |  |  |  | -30.45% | 37.17% | 35.24%
| Z-Score (Absolute) |  | 0.28 | 0.89 | 0.41 | 4.2 | 0.94 | 0.41 | 0.59 | 0.15 | 0.15 | 0.72 | 0.2 | 0.34 | 1.54 | 0.89 | 63.06 | 0.63 | 0.89 | 0.85
| Analytical results for Nd in OREAS 553b (Certified Value 20.2 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 1
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 1
| 1 | 1 | 22.353 | 27.1 | 37.4688 | 21.7148 | 18.775 | 15.5 | 15.9 | [19 empty] | 1
| 1 | 2 | 22.231 | 27.7 | 37.4328 | 22.568 | 18.94 | 15.7 | 17.09 | [19 empty] | 12
| 1 | 3 | 21.983 | 24.1 | 37.4584 | 23.9296 | 18.511 | 16.2 | 16.49 | [19 empty] | 16
| 1 | 4 | 23.359 | 26.4 | 37.4632 | 21.4306 | 18.649 | 16.2 | 15.31 | [19 empty] | 20.1566
| 1 | 5 | 22.773 | 23.6 | 37.4768 | 21.6407 | 18.809 | 15.5 | 15.29 | [19 empty] | 180
| 1 | 6 | 23.149 | 24.3 | 37.4184 | 21.7603 | 18.822 | 16 | 15.86
| Mean |  | 22.6413 | 25.5333 | 37.4531 | 22.174 | 18.751 | 15.85 | 15.99
| Median |  | 22.563 | 25.35 | 37.4608 | 21.7376 | 18.792 | 15.85 | 15.88
| Std Dev. |  | 0.54314 | 1.7443 | 0.022603 | 0.94409 | 0.15014 | 0.32711 | 0.6982
| Rel.Std.Dev. |  | 2.399% | 6.832% | 0.06035% | 4.258% | 0.8007% | 2.064% | 4.366%
| PDM3 |  | 12.33% | 26.67% | 85.81% | 10.01% | -6.973% | -21.37% | -20.67%
| Z-Score (Absolute) |  | 0.09 | 0.66 | 3.01 | 0 | 0.67 | 1.25 | 1.22
| Analytical results for Ni in OREAS 553b (Certified Value 115 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 0
| 1 | 1 | 133.3 | 111.3 | 121 | 118 | 115.45 | 143 | 114 | 124.6 | 117.9101 | 117 | 109 | 117.5 | 118.5 | 116 | 110.7023 | 111 | 116 | 127 | 134.5 | 114.5 | 86.31 | 110.21 | 91.2 | 119 | 114 |  | 1
| 1 | 2 | 129.8 | 116.3 | 121 | 119 | 111.62 | 144 | 115 | 127.3 | 119.0842 | 115 | 110.5 | 116 | 116 | 114.5 | 114.7057 | 109 | 123 | 133 | 130.1 | 114.1 | 85.92 | 114.718 | 96.4 | 115 | 118 |  | 29
| 1 | 3 | 131.5 | 116 | 115 | 118 | 114.48 | 145 | 114 | 126 | 123.8515 | 116.5 | 109 | 113 | 117 | 115 | 112.9668 | 114 | 122 | 131 | 132.6 | 112.8 | 87.84 | 112.367 | 98.8 | 118 | 114 |  | 16
| 1 | 4 | 132.1 | 113.1 | 123 | 118 | 112.37 | 150 | 114 | 122.8 | 120.4886 | 111.5 | 109.5 | 116 | 119.5 | 113 | 112.0086 | 112 | 124 | 132 | 131.8 | 111.8 | 84.16 | 112.741 | 99 | 117 | 112 |  | 115.4174
| 1 | 5 | 132.2 | 108 | 115 | 119 | 115.49 | 150 | 114 | 126.7 | 117.7673 | 110.5 | 108.5 | 117.5 | 117.5 | 109 | 110.5235 | 111 | 120 | 134 | 127.5 | 111.6 | 86.61 | 112.113 | 93.5 | 114 | 112 |  | 181
| 1 | 6 | 133.1 | 108.6 | 115 | 120 | 113.14 | 152 | 114 | 123.4 | 123.5102 | 114 | 107.5 | 116 | 119.5 | 112.5 | 108.7402 | 111 | 123 | 132 | 131.9 | 112.4 | 85.01 | 112.821 | 103 | 112 | 116
| Mean |  | 132 | 112.2167 | 118.3333 | 118.6667 | 113.7583 | 147.3333 | 114.1667 | 125.1333 | 120.4353 | 114.0833 | 109 | 116 | 118 | 113.3333 | 111.6079 | 111.3333 | 121.3333 | 131.5 | 131.4 | 112.8667 | 85.975 | 112.495 | 96.9833 | 115.8333 | 114.3333
| Median |  | 132.15 | 112.2 | 118 | 118.5 | 113.81 | 147.5 | 114 | 125.3 | 119.7864 | 114.5 | 109 | 116 | 118 | 113.75 | 111.3555 | 111 | 122.5 | 132 | 131.85 | 112.6 | 86.115 | 112.554 | 97.6 | 116 | 114
| Std Dev. |  | 1.2681 | 3.5628 | 3.7238 | 0.8165 | 1.6287 | 3.7771 | 0.40825 | 1.8239 | 2.7 | 2.6347 | 1 | 1.6432 | 1.4142 | 2.4833 | 2.0892 | 1.633 | 2.9439 | 2.429 | 2.3816 | 1.1961 | 1.2824 | 1.4486 | 4.2306 | 2.6394 | 2.3381
| Rel.Std.Dev. |  | 0.9607% | 3.175% | 3.147% | 0.6881% | 1.432% | 2.564% | 0.3576% | 1.458% | 2.242% | 2.309% | 0.9174% | 1.417% | 1.198% | 2.191% | 1.872% | 1.467% | 2.426% | 1.847% | 1.812% | 1.06% | 1.492% | 1.288% | 4.362% | 2.279% | 2.045%
| PDM3 |  | 14.37% | -2.773% | 2.526% | 2.815% | -1.437% | 27.65% | -1.084% | 8.418% | 4.348% | -1.156% | -5.56% | 0.5048% | 2.238% | -1.806% | -3.301% | -3.538% | 5.126% | 13.93% | 13.85% | -2.21% | -25.51% | -2.532% | -15.97% | 0.3604% | -0.9392%
| Z-Score (Absolute) |  | 3.25 | 0.39 | 0.74 | 0.8 | 0.11 | 6.07 | 0.03 | 1.99 | 1.12 | 0.05 | 0.98 | 0.31 | 0.67 | 0.18 | 0.5 | 0.55 | 1.29 | 3.16 | 3.14 | 0.27 | 5.22 | 0.34 | 3.19 | 0.28 | 0
| Analytical results for P in OREAS 553b (Certified Value 0.021 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 0.0233 | 0.02 | 0.036 | 0.0224 | 0.019 | 0.0198 | 0.0214 | 0.021147 | 0.021 | 0.021 | 0.019 | 0.023 | 0.02 | 0.021719 | 0.0192 | 0.02 | 0.0242 | 0.02 | 0.022 | 0.02 | 0.017041 | < 0.0009 | 0.0254 | 0.021 |  |  | 1
| 1 | 2 | 0.0231 | 0.02 | 0.036 | 0.0229 | 0.016 | 0.0197 | 0.0204 | 0.021869 | 0.021 | 0.021 | 0.019 | 0.022 | 0.02 | 0.021643 | 0.019 | 0.02 | 0.0241 | 0.019 | 0.022 | 0.02 | 0.018733 | < 0.0009 | 0.0276 | 0.021 |  |  | #N/A
| 1 | 3 | 0.0237 | 0.02 | 0.028 | 0.0227 | 0.02 | 0.02 | 0.0211 | 0.021707 | 0.021 | 0.021 | 0.019 | 0.022 | 0.021 | 0.021488 | 0.0189 | 0.02 | 0.0242 | 0.019 | 0.022 | 0.02 | 0.016966 | < 0.0009 | 0.0295 | 0.021 |  |  | 16
| 1 | 4 | 0.0239 | 0.02 | 0.037 | 0.0229 | 0.02 | 0.0197 | 0.0217 | 0.02227 | 0.021 | 0.021 | 0.02 | 0.023 | 0.02 | 0.021496 | 0.0191 | 0.02 | 0.0247 | 0.019 | 0.021 | 0.02 | 0.018728 | < 0.0009 | 0.0319 | 0.02 |  |  | 0.020547
| 1 | 5 | 0.0228 | 0.019 | 0.028 | 0.0222 | 0.02 | 0.0193 | 0.0215 | 0.020715 | 0.02 | 0.02 | 0.019 | 0.023 | 0.02 | 0.02134 | 0.0192 | 0.02 | 0.0262 | 0.019 | 0.021 | 0.02 | 0.018084 | < 0.0009 | 0.0303 | 0.02 |  |  | 182
| 1 | 6 | 0.024 | 0.019 | 0.03 | 0.0222 | 0.016 | 0.0194 | 0.0202 | 0.020237 | 0.021 | 0.02 | 0.02 | 0.022 | 0.02 | 0.021471 | 0.0192 | 0.02 | 0.0263 | 0.019 | 0.021 | 0.02 | 0.018551 | < 0.0009 | 0.0298 | 0.021
| Mean |  | 0.023467 | 0.019667 | 0.0325 | 0.02255 | 0.0185 | 0.01965 | 0.02105 | 0.021324 | 0.020833 | 0.020667 | 0.019333 | 0.0225 | 0.020167 | 0.021526 | 0.0191 | 0.02 | 0.02495 | 0.019167 | 0.0215 | 0.02 | 0.018017 |  | 0.029083 | 0.020667
| Median |  | 0.0235 | 0.02 | 0.033 | 0.02255 | 0.0195 | 0.0197 | 0.02125 | 0.021427 | 0.021 | 0.021 | 0.019 | 0.0225 | 0.02 | 0.021492 | 0.01915 | 0.02 | 0.02445 | 0.019 | 0.0215 | 0.02 | 0.018317 |  | 0.02965 | 0.021
| Std Dev. |  | 0.0004761 | 0.0005164 | 0.0042778 | 0.00032711 | 0.0019748 | 0.00025884 | 0.00061563 | 0.00076414 | 0.00040825 | 0.0005164 | 0.0005164 | 0.00054772 | 0.00040825 | 0.00013476 | 0.00012649 | 0 | 0.0010291 | 0.00040825 | 0.00054772 | 0 | 0.00082038 |  | 0.0022746 | 0.0005164
| Rel.Std.Dev. |  | 2.029% | 2.626% | 13.16% | 1.451% | 10.67% | 1.317% | 2.925% | 3.583% | 1.96% | 2.499% | 2.671% | 2.434% | 2.024% | 0.626% | 0.6623% | 0% | 4.125% | 2.13% | 2.548% | 0% | 4.553% |  | 7.821% | 2.499%
| PDM3 |  | 14.21% | -4.283% | 58.18% | 9.75% | -9.961% | -4.364% | 2.449% | 3.784% | 1.395% | 0.5838% | -5.905% | 9.507% | -1.85% | 4.768% | -7.041% | -2.661% | 21.43% | -6.717% | 4.64% | -2.661% | -12.31% |  | 41.55% | 0.5838%
| Z-Score (Absolute) |  | 1.61 | 0.57 | 6.79 | 1.08 | 1.24 | 0.58 | 0.22 | 0.38 | 0.1 | 0 | 0.77 | 1.05 | 0.29 | 0.49 | 0.9 | 0.38 | 2.46 | 0.86 | 0.48 | 0.38 | 1.52 | 11.61 | 4.83 | 0
| Analytical results for Pb in OREAS 553b (Certified Value 8.42 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 3
| 1 | 1 | 8.7 | 8.42 | 9.3 | 9.2 | 10 | 8.7 | 7.3 | 9.0288 | 8.5 | 8.1 | 8.2 | 8.6 | 8.5 | 9.967 | 8.2 | 8 | 9.7 | 11 | 8.5 | 7.4 | 17 | 10.27 | 6.78 | 12.1 | 8 |  | 1
| 1 | 2 | 8.5 | 8.15 | 9.5 | 9.2 | 11 | 8.7 | 7.8 | 9.7752 | 8.6 | 8.1 | 8.2 | 8.2 | 8.4 | 10.0662 | 7.9 | 8.1 | 9.9 | 12 | 7.8 | 7.3 | 16.3 | 10.42 | 6.85 | 11.7 | 8 |  | 13
| 1 | 3 | 8.5 | 8.87 | 8.3 | 9.4 | 7 | 8.8 | 7.5 | 9.6389 | 8.5 | 8 | 7.9 | 8.5 | 8.4 | 10.0609 | 8.4 | 8.2 | 9.6 | 12 | 8 | 7.3 | 17.9 | 10.42 | 7.02 | 12 | 8 |  | 16
| 1 | 4 | 8.9 | 7.97 | 9.5 | 9.3 | 12 | 8.7 | 8 | 9.1002 | 8 | 8.1 | 8 | 8.5 | 8 | 9.8213 | 8 | 8.2 | 9.9 | 10 | 8.1 | 7.5 | 16.2 | 11.64 | 7.02 | 12 | 8 |  | 8.42
| 1 | 5 | 8.7 | 7.34 | 8.3 | 9.3 | 7 | 8.8 | 7.4 | 8.8395 | 8 | 7.8 | 8.4 | 8.4 | 8.2 | 9.3466 | 8.1 | 8.1 | 10 | 11 | 7.9 | 7.2 | 17.7 | 10.19 | 6.7 | 11.9 | 8 |  | 183
| 1 | 6 | 9 | 7.55 | 8.5 | 9.2 | 10 | 8.9 | 7.9 | 8.5274 | 8.3 | 7.8 | 8.1 | 8.9 | 8.3 | 9.2321 | 7.8 | 8 | 10 | 12 | 8.2 | 7.3 | 16.8 | 11.32 | 6.95 | 11.6 | 8
| Mean |  | 8.7167 | 8.05 | 8.9 | 9.2667 | 9.5 | 8.7667 | 7.65 | 9.1517 | 8.3167 | 7.9833 | 8.1333 | 8.5167 | 8.3 | 9.749 | 8.0667 | 8.1 | 9.85 | 11.3333 | 8.0833 | 7.3333 | 16.9833 | 10.71 | 6.8867 | 11.8833 | 8
| Median |  | 8.7 | 8.06 | 8.9 | 9.25 | 10 | 8.75 | 7.65 | 9.0645 | 8.4 | 8.05 | 8.15 | 8.5 | 8.35 | 9.8941 | 8.05 | 8.1 | 9.9 | 11.5 | 8.05 | 7.3 | 16.9 | 10.42 | 6.9 | 11.95 | 8
| Std Dev. |  | 0.20412 | 0.5621 | 0.5933 | 0.08165 | 2.0736 | 0.08165 | 0.2881 | 0.47563 | 0.26394 | 0.1472 | 0.17512 | 0.23166 | 0.17889 | 0.36874 | 0.21602 | 0.089443 | 0.16432 | 0.8165 | 0.24833 | 0.10328 | 0.70261 | 0.61142 | 0.13201 | 0.19408 | 0
| Rel.Std.Dev. |  | 2.342% | 6.983% | 6.666% | 0.8811% | 21.83% | 0.9314% | 3.766% | 5.197% | 3.174% | 1.844% | 2.153% | 2.72% | 2.155% | 3.782% | 2.678% | 1.104% | 1.668% | 7.204% | 3.072% | 1.408% | 4.137% | 5.709% | 1.917% | 1.633% | 0%
| PDM3 |  | 3.523% | -4.395% | 5.7% | 10.06% | 12.83% | 4.117% | -9.145% | 8.689% | -1.228% | -5.186% | -3.405% | 1.148% | -1.426% | 15.78% | -4.197% | -3.801% | 16.98% | 34.6% | -3.999% | -12.91% | 101.7% | 27.2% | -18.21% | 41.13% | -4.988%
| Z-Score (Absolute) |  | 0.25 | 0.59 | 0.48 | 0.95 | 1.24 | 0.32 | 1.1 | 0.8 | 0.25 | 0.67 | 0.48 | 0 | 0.27 | 1.56 | 0.57 | 0.53 | 1.69 | 3.56 | 0.55 | 1.5 | 10.7 | 2.77 | 2.06 | 4.26 | 0.65
| Analytical results for Pd in OREAS 553b (Indicative Value < 10 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | Lab | [25 empty] | 1
| No | No | 09 | [25 empty] | ppb
|  |  | AR*OES/MS | [25 empty] | 1
|  |  | 01g | [25 empty] | 1
| 1 | 1 | < 10 | [25 empty] | 1
| 1 | 2 | < 10 | [25 empty] | 27
| 1 | 3 | < 10 | [25 empty] | 16
| 1 | 4 | < 10 | [25 empty] | < 10
| 1 | 5 | < 10 | [25 empty] | 33
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 553b (Certified Value 5.28 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 5.782 | 6.9 | 10.26 | 5.7687 | 4.748 | 4.01 | 4.44 | [19 empty] | 1
| 1 | 2 | 5.722 | 7 | 10.2944 | 5.9096 | 4.924 | 4.03 | 4.86 | [19 empty] | 14
| 1 | 3 | 5.625 | 6.3 | 10.2544 | 6.2829 | 4.744 | 4.15 | 4.59 | [19 empty] | 16
| 1 | 4 | 5.972 | 6.7 | 10.248 | 5.6624 | 4.821 | 4.17 | 4.39 | [19 empty] | 5.2784
| 1 | 5 | 5.851 | 6.1 | 10.2064 | 5.7333 | 4.821 | 3.99 | 4.41 | [19 empty] | 184
| 1 | 6 | 5.905 | 6.4 | 10.2312 | 5.6895 | 4.852 | 4.12 | 4.65
| Mean |  | 5.8095 | 6.5667 | 10.2491 | 5.8411 | 4.8183 | 4.0783 | 4.5567
| Median |  | 5.8165 | 6.55 | 10.2512 | 5.751 | 4.821 | 4.075 | 4.515
| Std Dev. |  | 0.12625 | 0.3559 | 0.029475 | 0.23312 | 0.067506 | 0.077567 | 0.18151
| Rel.Std.Dev. |  | 2.173% | 5.42% | 0.2876% | 3.991% | 1.401% | 1.902% | 3.983%
| PDM3 |  | 10.06% | 24.41% | 94.17% | 10.66% | -8.716% | -22.74% | -13.67%
| Z-Score (Absolute) |  | 0 | 0.52 | 3.02 | 0.02 | 0.67 | 1.18 | 0.85
| Analytical results for Pt in OREAS 553b (Indicative Value 3 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 09 | 25 | [24 empty] | ppb
|  |  | AR*OES/MS | AR*OES/MS | [24 empty] | 2
|  |  | 01g | 0.1g | [24 empty] | 2
| 1 | 1 | < 5 | 3 | [24 empty] | 1
| 1 | 2 | < 5 | 3 | [24 empty] | 2
| 1 | 3 | < 5 | 3 | [24 empty] | 16
| 1 | 4 | < 5 | 2 | [24 empty] | 3
| 1 | 5 | < 5 | 4 | [24 empty] | 31
| 1 | 6 | < 5 | 3
| Mean |  |  | 3
| Median |  |  | 3
| Std Dev. |  |  | 0.63246
| Rel.Std.Dev. |  |  | 21.08%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 553b (Certified Value 6.29 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 2
| 1 | 1 | 7.45 | 8.6 | 12.5 | 7 | 5.9 | 6.4785 | 5.6 | 5.5 | 5.7 | 7.2 | 6 | 8.0657 | 5.7 | 5.9 | 6.4 | 7.6 | 7.278 | 3.9 | 24.4 | 4.4 | [6 empty] | 1
| 1 | 2 | 7.51 | 8.8 | 12.6 | 7.6 | 5.6 | 7.1946 | 5.4 | 5.6 | 5.7 | 6.6 | 5.9 | 7.5302 | 5.5 | 6.2 | 6.3 | 7.2 | 7.602 | 3.95 | 23.9 | 4.51 | [6 empty] | 15
| 1 | 3 | 7.46 | 8.2 | 12.2 | 6.7 | 5.3 | 6.6632 | 5.4 | 5.6 | 5.5 | 6.8 | 5.9 | 7.7072 | 6.1 | 6 | 6.4 | 7.2 | 7.184 | 4.01 | 26.1 | 4.3 | [6 empty] | 16
| 1 | 4 | 7.43 | 7.9 | 12.1 | 7.4 | 5.9 | 7.3119 | 5.1 | 5.6 | 5.6 | 6.8 | 5.9 | 7.6771 | 5.7 | 6.2 | 6.4 | 7.2 | 7.474 | 3.99 | 24.8 | 4.29 | [6 empty] | 6.2944
| 1 | 5 | 7.27 | 7.3 | 12.8 | 7.1 | 5.4 | 6.5191 | 5.2 | 5.5 | 5.4 | 7 | 5.6 | 7.7669 | 5.8 | 6 | 6.3 | 7.5 | 7.467 | 3.93 | 22.6 | 4.25 | [6 empty] | 185
| 1 | 6 | 7.6 | 8.1 | 11.9 | 7.4 | 6 | 6.5915 | 5.3 | 5.3 | 5.7 | 6.8 | 5.8 | 7.2731 | 5.5 | 6.2 | 6.4 | 6.8 | 7.557 | 4.01 | 21.4 | 4.49
| Mean |  | 7.4533 | 8.15 | 12.35 | 7.2 | 5.6833 | 6.7931 | 5.3333 | 5.5167 | 5.6 | 6.8667 | 5.85 | 7.67 | 5.7167 | 6.0833 | 6.3667 | 7.25 | 7.427 | 3.965 | 23.8667 | 4.3733
| Median |  | 7.455 | 8.15 | 12.35 | 7.25 | 5.75 | 6.6273 | 5.35 | 5.55 | 5.65 | 6.8 | 5.9 | 7.6921 | 5.7 | 6.1 | 6.4 | 7.2 | 7.4705 | 3.97 | 24.15 | 4.35
| Std Dev. |  | 0.10857 | 0.53198 | 0.33912 | 0.32863 | 0.29269 | 0.36386 | 0.17512 | 0.1169 | 0.12649 | 0.20656 | 0.13784 | 0.26251 | 0.22286 | 0.13292 | 0.05164 | 0.28107 | 0.16284 | 0.045497 | 1.6633 | 0.11003
| Rel.Std.Dev. |  | 1.457% | 6.527% | 2.746% | 4.564% | 5.15% | 5.356% | 3.283% | 2.119% | 2.259% | 3.008% | 2.356% | 3.422% | 3.898% | 2.185% | 0.8111% | 3.877% | 2.193% | 1.147% | 6.969% | 2.516%
| PDM3 |  | 18.41% | 29.48% | 96.21% | 14.39% | -9.708% | 7.924% | -15.27% | -12.36% | -11.03% | 9.092% | -7.06% | 21.86% | -9.178% | -3.353% | 1.149% | 15.18% | 17.99% | -37.01% | 279.2% | -30.52%
| Z-Score (Absolute) |  | 0.67 | 1.2 | 4.4 | 0.47 | 0.68 | 0.16 | 0.95 | 0.81 | 0.75 | 0.22 | 0.56 | 0.83 | 0.66 | 0.38 | 0.16 | 0.51 | 0.65 | 1.99 | 13.17 | 1.68
| Analytical results for Re in OREAS 553b (Certified Value 0.026 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 21 | 25 | 26 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.1g | 0.4g | [11 empty] | 3
| 1 | 1 | 0.027 | 0.017 | < 0.05 | 0.025 | 0.026 | < 0.05 | 0.027 | 0.025 | 0.026 | 0.027 | 0.025 | 0.023047 | 0.024 | 0.026 | < 0.05 | [11 empty] | 1
| 1 | 2 | 0.024 | 0.018 | < 0.05 | 0.024 | 0.027 | < 0.05 | 0.028 | 0.026 | 0.023 | 0.026 | 0.024 | 0.028896 | 0.023 | 0.027 | < 0.05 | [11 empty] | 16
| 1 | 3 | 0.028 | 0.018 | < 0.05 | 0.029 | 0.026 | < 0.05 | 0.028 | 0.028 | 0.023 | 0.026 | 0.027 | 0.026521 | 0.022 | 0.027 | < 0.05 | [11 empty] | 16
| 1 | 4 | 0.027 | 0.019 | < 0.05 | 0.025 | 0.025 | < 0.05 | 0.026 | 0.026 | 0.025 | 0.026 | 0.025 | 0.025617 | 0.022 | 0.025 | < 0.05 | [11 empty] | 0.025986
| 1 | 5 | 0.024 | 0.017 | < 0.05 | 0.032 | 0.026 | < 0.05 | 0.023 | 0.026 | 0.025 | 0.026 | 0.025 | 0.023887 | 0.023 | 0.025 | < 0.05 | [11 empty] | 186
| 1 | 6 | 0.029 | 0.018 | < 0.05 | 0.025 | 0.027 | < 0.05 | 0.023 | 0.026 | 0.023 | 0.027 | 0.027 | 0.026702 | 0.022 | 0.026 | < 0.05
| Mean |  | 0.0265 | 0.017833 |  | 0.026667 | 0.026167 |  | 0.025833 | 0.026167 | 0.024167 | 0.026333 | 0.0255 | 0.025778 | 0.022667 | 0.026
| Median |  | 0.027 | 0.018 |  | 0.025 | 0.026 |  | 0.0265 | 0.026 | 0.024 | 0.026 | 0.025 | 0.026069 | 0.0225 | 0.026
| Std Dev. |  | 0.0020736 | 0.00075277 |  | 0.0031411 | 0.00075277 |  | 0.0023166 | 0.00098319 | 0.0013292 | 0.0005164 | 0.0012247 | 0.0021066 | 0.0008165 | 0.00089443
| Rel.Std.Dev. |  | 7.825% | 4.221% |  | 11.78% | 2.877% |  | 8.968% | 3.757% | 5.5% | 1.961% | 4.803% | 8.172% | 3.602% | 3.44%
| PDM3 |  | 1.976% | -31.37% |  | 2.617% | 0.6933% |  | -0.5894% | 0.6933% | -7.003% | 1.335% | -1.872% | -0.8007% | -12.78% | 0.05199%
| Z-Score (Absolute) |  | 0.67 | 7.42 | 0.73 | 0.83 | 0.36 | 0.73 | 0.05 | 0.36 | 1.51 | 0.52 | 0.26 | 0 | 2.91 | 0.21 | 0.73
| Analytical results for S in OREAS 553b (Certified Value 3.26 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | 3
| 1 | 1 | 3.4892 | 3.09 | 3.16 | 3.25 | 3.3477 | 3.77 | 3.27 | 3.38 | 3.2824 | 3.27 | 3.41 | 3.17 | 3.61 | 3.19 | 3.2972 | 3.34 | 3.26 | 3.17 | 3.2455 | 3.24 | 3.27 | 3.43 | 3.095 | 3.33 | 3.35 | 3.12 | 1
| 1 | 2 | 3.6251 | 3.21 | 3.224 | 3.32 | 3.352 | 3.88 | 3.32 | 3.42 | 3.2509 | 3.23 | 3.42 | 3.12 | 3.54 | 3.15 | 3.3023 | 3.26 | 3.32 | 3.32 | 3.2197 | 3.13 | 3.26 | 3.41 | 3.088 | 3.3 | 3.15 | 3.08 | 31
| 1 | 3 | 3.5551 | 3.16 | 3.317 | 3.3 | 3.3497 | 3.94 | 3.31 | 3.48 | 3.3108 | 3.27 | 3.35 | 3.12 | 3.52 | 3.2 | 3.2704 | 3.29 | 3.31 | 3.33 | 3.2468 | 3.08 | 3.2 | 3.56 | 3.105 | 3.34 | 3.17 | 3.11 | 16
| 1 | 4 | 3.5288 | 3.09 | 3.319 | 3.33 | 3.354 | 4 | 3.31 | 3.37 | 3.2609 | 3.17 | 3.38 | 3.16 | 3.62 | 3.12 | 3.2587 | 3.26 | 3.34 | 3.29 | 3.3166 | 3.19 | 3.23 | 3.41 | 3.09 | 3.3 | 3.4 | 3.12 | 3.2631
| 1 | 5 | 3.6593 | 3.02 | 3.358 | 3.27 | 3.3603 | 4.02 | 3.31 | 3.45 | 3.2519 | 3.16 | 3.33 | 3.24 | 3.56 | 3.1 | 3.2941 | 3.3 | 3.28 | 3.33 | 3.1133 | 3.22 | 3.23 | 3.45 | 3.132 | 3.28 | 3.41 | 3.08 | 187
| 1 | 6 | 3.6341 | 2.99 | 3.428 | 3.38 | 3.3579 | 4.07 | 3.31 | 3.43 | 3.2975 | 3.23 | 3.33 | 3.24 | 3.55 | 3.12 | 3.2639 | 3.27 | 3.34 | 3.34 | 3.2479 | 3.19 | 3.22 | 3.42 | 3.134 | 3.37 | 3.38 | 3.11
| Mean |  | 3.5819 | 3.0933 | 3.301 | 3.3083 | 3.3536 | 3.9467 | 3.305 | 3.4217 | 3.2757 | 3.2217 | 3.37 | 3.175 | 3.5667 | 3.1467 | 3.2811 | 3.2867 | 3.3083 | 3.2967 | 3.2316 | 3.175 | 3.235 | 3.4467 | 3.1073 | 3.32 | 3.31 | 3.1033
| Median |  | 3.5901 | 3.09 | 3.318 | 3.31 | 3.353 | 3.97 | 3.31 | 3.425 | 3.2716 | 3.23 | 3.365 | 3.165 | 3.555 | 3.135 | 3.2822 | 3.28 | 3.315 | 3.325 | 3.2461 | 3.19 | 3.23 | 3.425 | 3.1 | 3.315 | 3.365 | 3.11
| Std Dev. |  | 0.067399 | 0.082624 | 0.095674 | 0.046224 | 0.0048191 | 0.10875 | 0.017607 | 0.041673 | 0.025108 | 0.047504 | 0.039497 | 0.054314 | 0.039833 | 0.040825 | 0.018934 | 0.030768 | 0.032506 | 0.064395 | 0.06641 | 0.059582 | 0.025884 | 0.057504 | 0.020743 | 0.032863 | 0.11815 | 0.018619
| Rel.Std.Dev. |  | 1.882% | 2.671% | 2.898% | 1.397% | 0.1437% | 2.756% | 0.5327% | 1.218% | 0.7665% | 1.475% | 1.172% | 1.711% | 1.117% | 1.297% | 0.5771% | 0.9361% | 0.9826% | 1.953% | 2.055% | 1.877% | 0.8001% | 1.668% | 0.6675% | 0.9899% | 3.57% | 0.6%
| PDM3 |  | 9.771% | -5.202% | 1.162% | 1.387% | 2.774% | 20.95% | 1.284% | 4.86% | 0.3873% | -1.269% | 3.276% | -2.7% | 9.303% | -3.568% | 0.552% | 0.7225% | 1.387% | 1.029% | -0.964% | -2.7% | -0.8608% | 5.626% | -4.773% | 1.744% | 1.438% | -4.896%
| Z-Score (Absolute) |  | 2.76 | 2 | 0.02 | 0.09 | 0.53 | 6.31 | 0.06 | 1.2 | 0.23 | 0.75 | 0.69 | 1.21 | 2.61 | 1.48 | 0.17 | 0.12 | 0.09 | 0.02 | 0.65 | 1.21 | 0.62 | 1.44 | 1.87 | 0.21 | 0.11 | 1.91
| Analytical results for Sb in OREAS 553b (Certified Value 1.86 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 2
| 1 | 1 | 2.66 | 2.22 | 2.03 | 1.78 | < 5 | 1.56 | 1.0246 | 1.86 | 2.01 | 1.88 | 1.77 | 1.92 | 1.9772 | 1.81 | 1.88 | 1.81 | 1.58 | 1.5 | 1.96 | < 5 | 1.978 | 0.48 | 1.4 | 1.91 |  |  | 1
| 1 | 2 | 2.49 | 2.5 | 2.09 | 1.77 | < 5 | 1.65 | 0.97362 | 1.84 | 2.06 | 1.94 | 1.65 | 1.94 | 2.1425 | 1.88 | 1.98 | 1.92 | 1.49 | 1.5 | 1.84 | < 5 | 2.084 | 0.53 | 1.38 | 1.91 |  |  | 17
| 1 | 3 | 2.86 | 2.28 | 2.29 | 1.94 | < 5 | 1.54 | 1.0389 | 1.82 | 2.06 | 1.85 | 1.68 | 1.88 | 2.0557 | 2.04 | 2 | 1.91 | 1.58 | 1.6 | 1.9 | < 5 | 2.075 | 0.57 | 1.38 | 1.91 |  |  | 16
| 1 | 4 | 2.57 | 2.27 | 2.1 | 1.78 | < 5 | 1.52 | 1.0279 | 1.7 | 2.02 | 1.82 | 1.74 | 1.94 | 1.9975 | 1.94 | 2.05 | 1.81 | 1.57 | 1.6 | 1.95 | < 5 | 2.122 | 0.57 | 1.37 | 1.95 |  |  | 1.8554
| 1 | 5 | 2.5 | 2.17 | 2.34 | 1.93 | < 5 | 1.63 | 0.95443 | 1.8 | 1.92 | 1.82 | 1.7 | 1.78 | 2.0136 | 1.86 | 1.97 | 1.85 | 1.62 | 1.6 | 1.95 | < 5 | 2.086 | 0.6 | 1.41 | 1.9 |  |  | 188
| 1 | 6 | 2.71 | 2.12 | 2.43 | 1.87 | < 5 | 1.51 | 0.94205 | 1.71 | 1.82 | 1.85 | 1.61 | 1.81 | 2.0139 | 1.8 | 2 | 1.71 | 1.54 | 1.6 | 1.91 | < 5 | 2.12 | 0.68 | 1.27 | 1.91
| Mean |  | 2.6317 | 2.26 | 2.2133 | 1.845 |  | 1.5683 | 0.99359 | 1.7883 | 1.9817 | 1.86 | 1.6917 | 1.8783 | 2.0334 | 1.8883 | 1.98 | 1.835 | 1.5633 | 1.5667 | 1.9183 |  | 2.0775 | 0.57167 | 1.3683 | 1.915
| Median |  | 2.615 | 2.245 | 2.195 | 1.825 |  | 1.55 | 0.99913 | 1.81 | 2.015 | 1.85 | 1.69 | 1.9 | 2.0137 | 1.87 | 1.99 | 1.83 | 1.575 | 1.6 | 1.93 |  | 2.085 | 0.57 | 1.38 | 1.91
| Std Dev. |  | 0.14162 | 0.13221 | 0.16158 | 0.078677 |  | 0.058452 | 0.041907 | 0.067651 | 0.094322 | 0.045166 | 0.058452 | 0.068823 | 0.059367 | 0.089981 | 0.056214 | 0.077395 | 0.044121 | 0.05164 | 0.04535 |  | 0.05255 | 0.067355 | 0.050365 | 0.017607
| Rel.Std.Dev. |  | 5.381% | 5.85% | 7.3% | 4.264% |  | 3.727% | 4.218% | 3.783% | 4.76% | 2.428% | 3.455% | 3.664% | 2.92% | 4.765% | 2.839% | 4.218% | 2.822% | 3.296% | 2.364% |  | 2.529% | 11.78% | 3.681% | 0.9194%
| PDM3 |  | 41.84% | 21.81% | 19.29% | -0.5595% |  | -15.47% | -46.45% | -3.614% | 6.806% | 0.2489% | -8.824% | 1.237% | 9.595% | 1.776% | 6.717% | -1.098% | -15.74% | -15.56% | 3.393% |  | 11.97% | -69.19% | -26.25% | 3.213%
| Z-Score (Absolute) |  | 2.62 | 1.32 | 1.15 | 0.13 | 2.16 | 1.1 | 3.11 | 0.33 | 0.34 | 0.08 | 0.67 | 0.02 | 0.52 | 0.02 | 0.34 | 0.17 | 1.12 | 1.11 | 0.12 | 2.16 | 0.68 | 4.58 | 1.8 | 0.11
| Analytical results for Sc in OREAS 553b (Certified Value 3.3 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 2
| 1 | 1 | 3.3 | 3.2 | 3.5 | 4 | 3.8 | 3.7 | 3.4 | 3.2454 | 3.3 | 3.1 | 3.3 | 3.6 | 3.1 | 3.2176 | 3.1 | 3.4 | 4 | 3.2 | 3.5 | 3.2 | 2.903 | 2.7 | 5 | 3 |  |  | 1
| 1 | 2 | 3.6 | 3.6 | 3.3 | 4 | 3.7 | 3.5 | 3.7 | 3.4957 | 3.2 | 3.1 | 3.3 | 3.3 | 3.3 | 3.4803 | 2.7 | 3.6 | 4 | 2.8 | 3.2 | 3 | 3.005 | 2.6 | 4 | 2.8 |  |  | 18
| 1 | 3 | 3.5 | 3.2 | 3.2 | 4 | 3.7 | 3.7 | 3.3 | 3.5556 | 3.3 | 3.2 | 3.1 | 3.5 | 3.1 | 3.5702 | 3 | 3.3 | 4 | 3 | 3.4 | 3.2 | 2.928 | 2.7 | 5 | 2.9 |  |  | 16
| 1 | 4 | 3.6 | 3.3 | 3.3 | 4 | 3.9 | 3.5 | 3.6 | 3.6043 | 3 | 3.1 | 3.3 | 3.5 | 3.1 | 3.3994 | 2.9 | 3.6 | 4 | 2.9 | 3.2 | 3 | 2.999 | 2.8 | 6 | 2.9 |  |  | 3.2999
| 1 | 5 | 3.5 | 3.1 | 2.9 | 4 | 4 | 3.6 | 3.3 | 3.2388 | 3 | 3 | 3.2 | 3.5 | 3 | 3.3556 | 3 | 3.4 | 4 | 2.9 | 3.2 | 3.1 | 2.747 | 2.5 | 5 | 2.7 |  |  | 189
| 1 | 6 | 3.5 | 3.2 | 3.2 | 4 | 3.8 | 3.5 | 3.7 | 3.4327 | 3.1 | 3.1 | 3.3 | 3.3 | 3.1 | 3.3847 | 2.7 | 3.4 | 4 | 2.9 | 3.2 | 3 | 2.821 | 2.7 | 5 | 2.9
| Mean |  | 3.5 | 3.2667 | 3.2333 | 4 | 3.8167 | 3.5833 | 3.5 | 3.4288 | 3.15 | 3.1 | 3.25 | 3.45 | 3.1167 | 3.4013 | 2.9 | 3.45 | 4 | 2.95 | 3.2833 | 3.0833 | 2.9005 | 2.6667 | 5 | 2.8667
| Median |  | 3.5 | 3.2 | 3.25 | 4 | 3.8 | 3.55 | 3.5 | 3.4642 | 3.15 | 3.1 | 3.3 | 3.5 | 3.1 | 3.3921 | 2.95 | 3.4 | 4 | 2.9 | 3.2 | 3.05 | 2.9155 | 2.7 | 5 | 2.9
| Std Dev. |  | 0.10954 | 0.17512 | 0.19664 | 0 | 0.1169 | 0.098319 | 0.18974 | 0.15564 | 0.13784 | 0.063246 | 0.083666 | 0.12247 | 0.098319 | 0.11907 | 0.16733 | 0.12247 | 0 | 0.13784 | 0.13292 | 0.098319 | 0.10126 | 0.10328 | 0.63246 | 0.10328
| Rel.Std.Dev. |  | 3.13% | 5.361% | 6.082% | 0% | 3.063% | 2.744% | 5.421% | 4.539% | 4.376% | 2.04% | 2.574% | 3.55% | 3.155% | 3.501% | 5.77% | 3.55% | 0% | 4.673% | 4.048% | 3.189% | 3.491% | 3.873% | 12.65% | 3.603%
| PDM3 |  | 6.064% | -1.007% | -2.017% | 21.22% | 15.66% | 8.59% | 6.064% | 3.906% | -4.542% | -6.057% | -1.512% | 4.549% | -5.552% | 3.074% | -12.12% | 4.549% | 21.22% | -10.6% | -0.5014% | -6.562% | -12.1% | -19.19% | 51.52% | -13.13%
| Z-Score (Absolute) |  | 0.73 | 0.03 | 0.13 | 2.35 | 1.75 | 1 | 0.73 | 0.5 | 0.4 | 0.57 | 0.08 | 0.57 | 0.51 | 0.41 | 1.21 | 0.57 | 2.35 | 1.05 | 0.03 | 0.62 | 1.21 | 1.97 | 5.58 | 1.32
| Analytical results for Se in OREAS 553b (Certified Value 5.65 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  |  | 3
| 1 | 1 | 6 | 2 | 5.6 | 5 | 6 | 5 | 5.3772 | 5.4 | 5.8 | 5.3 | 5.7 | 5.7 | 6.1229 | 6 | 7.28 | 5.6 | 6 | < 5 | 4.024 | < 0.1 | 4.95 | 4 |  |  |  |  | 1
| 1 | 2 | 6 | 2.1 | 5.8 | 6 | 6 | 7 | 5.2251 | 5.4 | 6.4 | 5.3 | 5.5 | 5.9 | 6.0457 | 6 | 7.03 | 5.2 | 5.6 | < 5 | 4.204 | < 0.1 | 5.11 | 4 |  |  |  |  | 21
| 1 | 3 | 6 | 4.9 | 5.2 | 6 | 6 | 6 | 5.4497 | 6.3 | 6.1 | 5.3 | 5.8 | 5.9 | 5.553 | 6 | 7.19 | 5.3 | 6 | < 5 | 3.56 | < 0.1 | 5.37 | 4 |  |  |  |  | 16
| 1 | 4 | 6 | 4.9 | 5.7 | 6 | 6 | 7 | 5.2742 | 5.3 | 5.9 | 5.7 | 5.7 | 6 | 5.7119 | 6 | 7.3 | 5.3 | 5.7 | < 5 | 4.569 | < 0.1 | 5.06 | 4 |  |  |  |  | 5.6527
| 1 | 5 | 6 | 4.9 | 5.1 | 6 | 6 | 6 | 5.5959 | 5.6 | 5.7 | 5.2 | 5.7 | 5.3 | 5.6748 | 6 | 7.29 | 5.5 | 6 | < 5 | 3.793 | < 0.1 | 4.85 | 4 |  |  |  |  | 190
| 1 | 6 | 6 | 4.3 | 5.2 | 6 | 6 | 5 | 5.6356 | 5.2 | 5.7 | 5.9 | 5.6 | 5.3 | 5.7673 | 5 | 7.37 | 5.4 | 6.1 | < 5 | 4.14 | < 0.1 | 4.99 | 4
| Mean |  | 6 | 3.85 | 5.4333 | 5.8333 | 6 | 6 | 5.4263 | 5.5333 | 5.9333 | 5.45 | 5.6667 | 5.6833 | 5.8126 | 5.8333 | 7.2433 | 5.3833 | 5.9 |  | 4.0483 |  | 5.055 | 4
| Median |  | 6 | 4.6 | 5.4 | 6 | 6 | 6 | 5.4134 | 5.4 | 5.85 | 5.3 | 5.7 | 5.8 | 5.7396 | 6 | 7.285 | 5.35 | 6 |  | 4.082 |  | 5.025 | 4
| Std Dev. |  | 0 | 1.4139 | 0.30111 | 0.40825 | 0 | 0.89443 | 0.16683 | 0.39833 | 0.27325 | 0.28107 | 0.10328 | 0.31252 | 0.22325 | 0.40825 | 0.11928 | 0.1472 | 0.2 |  | 0.34871 |  | 0.17863 | 0
| Rel.Std.Dev. |  | 0% | 36.72% | 5.542% | 6.999% | 0% | 14.91% | 3.075% | 7.199% | 4.605% | 5.157% | 1.823% | 5.499% | 3.841% | 6.999% | 1.647% | 2.734% | 3.39% |  | 8.614% |  | 3.534% | 0%
| PDM3 |  | 6.143% | -31.89% | -3.881% | 3.195% | 6.143% | 6.143% | -4.006% | -2.112% | 4.964% | -3.587% | 0.2463% | 0.5412% | 2.828% | 3.195% | 28.14% | -4.766% | 4.374% |  | -28.38% |  | -10.57% | -29.24%
| Z-Score (Absolute) |  | 0.85 | 3.73 | 0.35 | 0.5 | 0.85 | 0.85 | 0.37 | 0.14 | 0.71 | 0.32 | 0.14 | 0.18 | 0.45 | 0.5 | 3.5 | 0.46 | 0.64 | 6.6 | 3.3 | 11.82 | 1.16 | 3.41
| Analytical results for Sm in OREAS 553b (Certified Value 3.2 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | 24 | 26 | [19 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [19 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | 0.5g | 0.4g | [19 empty] | 2
| 1 | 1 | 3.466 | 4 | 5.2659 | 3.3882 | 2.932 | 2.51 | 2.92 | [19 empty] | 1
| 1 | 2 | 3.347 | 4.3 | 5.2138 | 3.6054 | 3.044 | 2.53 | 3.02 | [19 empty] | 5
| 1 | 3 | 3.356 | 3.8 | 5.2342 | 3.7262 | 2.881 | 2.65 | 2.88 | [19 empty] | 16
| 1 | 4 | 3.608 | 3.5 | 5.266 | 3.4534 | 2.918 | 2.64 | 2.78 | [19 empty] | 3.197
| 1 | 5 | 3.483 | 3.6 | 5.2485 | 3.4429 | 2.97 | 2.52 | 2.72 | [19 empty] | 191
| 1 | 6 | 3.633 | 3.7 | 5.2289 | 3.4568 | 2.93 | 2.6 | 2.78
| Mean |  | 3.4822 | 3.8167 | 5.2429 | 3.5121 | 2.9458 | 2.575 | 2.85
| Median |  | 3.4745 | 3.75 | 5.2414 | 3.4551 | 2.931 | 2.565 | 2.83
| Std Dev. |  | 0.12086 | 0.29269 | 0.021049 | 0.12743 | 0.05593 | 0.062849 | 0.11082
| Rel.Std.Dev. |  | 3.471% | 7.669% | 0.4015% | 3.628% | 1.899% | 2.441% | 3.888%
| PDM3 |  | 8.921% | 19.38% | 64% | 9.859% | -7.855% | -19.45% | -10.85%
| Z-Score (Absolute) |  | 0 | 0.42 | 2.21 | 0.04 | 0.67 | 1.14 | 0.79
| Analytical results for Sn in OREAS 553b (Certified Value 1.28 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [6 empty] | 3
| 1 | 1 | 1.36 | 1.16 | 1.6 | 1.2 | 1.3 | 1.2088 | 1.4 | 1.3 | 1.3 | 1.3 | 1.4 | 1.3789 | 1.22 | 1.1 | 1.3 | < 5 | 1.242 | 0.96 | 1.2 | 1.5 | [6 empty] | 1
| 1 | 2 | 1.34 | 1.2 | 1.7 | 1.2 | 1.4 | 1.2371 | 1.4 | 1.3 | 1.3 | 1.2 | 1.3 | 1.4071 | 1.21 | 1.2 | 1.3 | < 5 | 1.233 | 0.99 | 1.2 | 1.5 | [6 empty] | 22
| 1 | 3 | 1.32 | 1.25 | 1.9 | 1.3 | 1.3 | 1.2763 | 1.4 | 1.3 | 1.3 | 1.3 | 1.3 | 1.4064 | 1.39 | 1.2 | 1.3 | < 5 | 1.203 | 1 | 1.3 | 1.5 | [6 empty] | 16
| 1 | 4 | 1.33 | 1.2 | 1.9 | 1.2 | 1.3 | 1.2588 | 1.3 | 1.3 | 1.3 | 1.3 | 1.4 | 1.3534 | 1.21 | 1.2 | 1.3 | < 5 | 1.28 | 1.01 | 1.3 | 1.5 | [6 empty] | 1.2775
| 1 | 5 | 1.36 | 1.23 | 1.8 | 1.2 | 1.4 | 1.1903 | 1.3 | 1.2 | 1.3 | 1.3 | 1.3 | 1.3727 | 1.29 | 1.2 | 1.3 | < 5 | 1.217 | 0.96 | 1.2 | 1.5 | [6 empty] | 192
| 1 | 6 | 1.31 | 1.2 | 1.7 | 1.1 | 1.3 | 1.2157 | 1.3 | 1.2 | 1.3 | 1.3 | 1.3 | 1.3161 | 1.34 | 1.2 | 1.3 | < 5 | 1.22 | 0.99 | 1.2 | 1.4
| Mean |  | 1.3367 | 1.2067 | 1.7667 | 1.2 | 1.3333 | 1.2312 | 1.35 | 1.2667 | 1.3 | 1.2833 | 1.3333 | 1.3724 | 1.2767 | 1.1833 | 1.3 |  | 1.2325 | 0.985 | 1.2333 | 1.4833
| Median |  | 1.335 | 1.2 | 1.75 | 1.2 | 1.3 | 1.2264 | 1.35 | 1.3 | 1.3 | 1.3 | 1.3 | 1.3758 | 1.255 | 1.2 | 1.3 |  | 1.2265 | 0.99 | 1.2 | 1.5
| Std Dev. |  | 0.020656 | 0.030768 | 0.12111 | 0.063246 | 0.05164 | 0.032417 | 0.054772 | 0.05164 | 0 | 0.040825 | 0.05164 | 0.034443 | 0.076333 | 0.040825 | 0 |  | 0.026883 | 0.020736 | 0.05164 | 0.040825
| Rel.Std.Dev. |  | 1.545% | 2.55% | 6.855% | 5.27% | 3.873% | 2.633% | 4.057% | 4.077% | 0% | 3.181% | 3.873% | 2.51% | 5.979% | 3.45% | 0% |  | 2.181% | 2.105% | 4.187% | 2.752%
| PDM3 |  | 4.634% | -5.542% | 38.29% | -6.064% | 4.373% | -3.623% | 5.678% | -0.8453% | 1.764% | 0.4594% | 4.373% | 7.433% | -0.06246% | -7.369% | 1.764% |  | -3.52% | -22.89% | -3.455% | 16.12%
| Z-Score (Absolute) |  | 0.52 | 0.98 | 5.45 | 1.05 | 0.48 | 0.69 | 0.67 | 0.29 | 0.1 | 0.1 | 0.48 | 0.93 | 0.17 | 1.24 | 0.1 | 13.87 | 0.68 | 3.52 | 0.67 | 2.2
| Analytical results for Sr in OREAS 553b (Certified Value 16.4 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 1
| 1 | 1 | 18.19 | 15.2 | 22.6 | 21.7 | 16.1 | 16.4 | 14.1 | 16.8677 | 16.2 | 16 | 14.8 | 17.7 | 15.7 | 17.0294 | 12.5 | 16.2 | 28.1 | 15 | 16 | 18.6 | 14.473 | 21 | 24.7 | 13.4 |  |  | 1
| 1 | 2 | 18.02 | 16.6 | 23.5 | 21.6 | 17 | 15.8 | 13.9 | 18.0021 | 15.8 | 16.4 | 14.9 | 15.9 | 15.9 | 17.1148 | 12.3 | 16.4 | 29.5 | 13 | 16 | 18.6 | 15.107 | 21 | 21.4 | 14 |  |  | #N/A
| 1 | 3 | 17.79 | 16.7 | 19 | 21.6 | 16.6 | 15.1 | 12.7 | 17.1411 | 16.2 | 16.2 | 13.4 | 17.4 | 15.9 | 17.2858 | 12.7 | 16.5 | 27.6 | 14 | 16 | 18.8 | 14.399 | 21 | 25.9 | 13.4 |  |  | 16
| 1 | 4 | 17.75 | 15.4 | 22.9 | 21.8 | 16.4 | 15.6 | 14.3 | 18.1635 | 15 | 16.1 | 14.4 | 17.4 | 15.4 | 16.6545 | 12.9 | 16.8 | 28.6 | 13 | 15 | 19.3 | 14.566 | 20 | 25 | 12.8 |  |  | 16.4063
| 1 | 5 | 17.42 | 14.4 | 18.9 | 21 | 16.4 | 15.6 | 12.9 | 16.4823 | 15.2 | 15.6 | 14.8 | 16.9 | 14.8 | 16.794 | 12.8 | 16.4 | 28.9 | 14 | 15 | 18.2 | 14.637 | 20 | 22.8 | 13.3 |  |  | 193
| 1 | 6 | 17.78 | 15.7 | 20.2 | 21.1 | 16.6 | 15.7 | 13.2 | 17.0475 | 15.9 | 15.4 | 14.3 | 16.9 | 15.7 | 17.0365 | 12.6 | 16.6 | 29.2 | 14 | 15 | 18.8 | 14.575 | 22 | 23.5 | 13.9
| Mean |  | 17.825 | 15.6667 | 21.1833 | 21.4667 | 16.5167 | 15.7 | 13.5167 | 17.284 | 15.7167 | 15.95 | 14.4333 | 17.0333 | 15.5667 | 16.9858 | 12.6333 | 16.4833 | 28.65 | 13.8333 | 15.5 | 18.7167 | 14.6262 | 20.8333 | 23.8833 | 13.4667
| Median |  | 17.785 | 15.55 | 21.4 | 21.6 | 16.5 | 15.65 | 13.55 | 17.0943 | 15.85 | 16.05 | 14.6 | 17.15 | 15.7 | 17.033 | 12.65 | 16.45 | 28.75 | 14 | 15.5 | 18.7 | 14.5705 | 21 | 24.1 | 13.4
| Std Dev. |  | 0.26235 | 0.8756 | 2.0624 | 0.33267 | 0.29944 | 0.41952 | 0.67057 | 0.66053 | 0.50761 | 0.37815 | 0.56095 | 0.6377 | 0.41793 | 0.22702 | 0.21602 | 0.20412 | 0.7064 | 0.75277 | 0.54772 | 0.36009 | 0.25005 | 0.75277 | 1.6412 | 0.43665
| Rel.Std.Dev. |  | 1.472% | 5.589% | 9.736% | 1.55% | 1.813% | 2.672% | 4.961% | 3.822% | 3.23% | 2.371% | 3.887% | 3.744% | 2.685% | 1.337% | 1.71% | 1.238% | 2.466% | 5.442% | 3.534% | 1.924% | 1.71% | 3.613% | 6.872% | 3.242%
| PDM3 |  | 8.648% | -4.508% | 29.12% | 30.84% | 0.673% | -4.305% | -17.61% | 5.35% | -4.203% | -2.781% | -12.03% | 3.822% | -5.118% | 3.533% | -23% | 0.4698% | 74.63% | -15.68% | -5.524% | 14.08% | -10.85% | 26.98% | 45.57% | -17.92%
| Z-Score (Absolute) |  | 0.68 | 0.23 | 2.09 | 2.21 | 0.13 | 0.22 | 1.14 | 0.45 | 0.21 | 0.11 | 0.75 | 0.34 | 0.27 | 0.32 | 1.51 | 0.11 | 5.24 | 1 | 0.3 | 1.05 | 0.67 | 1.95 | 3.23 | 1.16
| Analytical results for Ta in OREAS 553b (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 24 | 25 | 26 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [8 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 7
| 1 | 3 | < 0.05 | < 0.05 | 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01 | [8 empty] | 194
| 1 | 6 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.3 | < 0.05 | < 0.1 | < 5 | < 0.01 | < 0.01 | < 0.01
| Mean |  |  |  | 0.01
| Median |  |  |  | 0.01
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 10.11 | 10.11 | 0.67 | 10.11 | 10.11 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 77.55 | 10.11 | 23.6 | 1345.25 | 0.67 | 0.67 | 0.67
| Analytical results for Tb in OREAS 553b (Certified Value 0.19 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 05 | 08 | 15 | 16 | 17 | 22 | 23 | 24 | 26 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | [15 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.4g | [15 empty] | 3
| 1 | 1 | 0.192 | 0.2 | 0.42368 | 0.17 | 0.20921 | 0.19 | 0.18 | < 5 | 0.173 | 0.14 | 0.21 | [15 empty] | 1
| 1 | 2 | 0.187 | 0.2 | 0.41928 | 0.2 | 0.22424 | 0.19 | 0.18 | < 5 | 0.181 | 0.15 | 0.21 | [15 empty] | 8
| 1 | 3 | 0.188 | 0.2 | 0.42544 | 0.18 | 0.2217 | 0.17 | 0.18 | < 5 | 0.174 | 0.15 | 0.21 | [15 empty] | 16
| 1 | 4 | 0.195 | 0.2 | 0.43152 | 0.17 | 0.20879 | 0.17 | 0.19 | < 5 | 0.178 | 0.15 | 0.2 | [15 empty] | 0.19025
| 1 | 5 | 0.196 | 0.2 | 0.4376 | 0.19 | 0.21001 | 0.17 | 0.18 | < 5 | 0.175 | 0.14 | 0.2 | [15 empty] | 195
| 1 | 6 | 0.197 | 0.2 | 0.42112 | 0.2 | 0.20675 | 0.19 | 0.18 | < 5 | 0.174 | 0.14 | 0.19
| Mean |  | 0.1925 | 0.2 | 0.42644 | 0.185 | 0.21345 | 0.18 | 0.18167 |  | 0.17583 | 0.145 | 0.20333
| Median |  | 0.1935 | 0.2 | 0.42456 | 0.185 | 0.20961 | 0.18 | 0.18 |  | 0.1745 | 0.145 | 0.205
| Std Dev. |  | 0.0042308 | 3.0405e-17 | 0.0069071 | 0.013784 | 0.0074954 | 0.010954 | 0.0040825 |  | 0.0030605 | 0.0054772 | 0.008165
| Rel.Std.Dev. |  | 2.198% | 1.52e-14% | 1.62% | 7.451% | 3.512% | 6.086% | 2.247% |  | 1.741% | 3.777% | 4.016%
| PDM3 |  | 1.18% | 5.122% | 124.1% | -2.762% | 12.19% | -5.39% | -4.514% |  | -7.58% | -23.79% | 6.874%
| Z-Score (Absolute) |  | 0.18 | NA | 11.66 | 0.18 | 1.21 | 0.43 | 0.35 | 113.35 | 0.63 | 2.15 | 0.72
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 553b (Certified Value 0.048 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 3
| 1 | 1 | 0.04 | 0.04 | < 0.05 | 0.1 | 0.06 | 0.051429 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | < 0.2 | < 0.05 | 0.76 | < 0.2 | < 0.1 | < 5 | 0.052 | 0.04 | 0.04 | < 0.05 | [5 empty] | 1
| 1 | 2 | 0.03 | 0.05 | 0.05 | 0.09 | < 0.05 | < 0.05 | 0.04 | 0.04 | 0.03 | 0.05 | 0.04 | < 0.2 | < 0.05 | 0.4 | < 0.2 | < 0.1 | < 5 | 0.066 | 0.04 | 0.05 | < 0.05 | [5 empty] | 23
| 1 | 3 | < 0.02 | 0.07 | < 0.05 | 0.09 | 0.06 | 0.070937 | 0.02 | 0.05 | 0.05 | 0.05 | 0.05 | < 0.2 | < 0.05 | 0.64 | < 0.2 | < 0.1 | < 5 | 0.04 | 0.04 | 0.04 | 0.06 | [5 empty] | 16
| 1 | 4 | < 0.02 | 0.07 | < 0.05 | 0.1 | < 0.05 | 0.065406 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | < 0.2 | < 0.05 | 0.44 | < 0.2 | < 0.1 | < 5 | 0.06 | 0.04 | 0.05 | 0.06 | [5 empty] | 0.047667
| 1 | 5 | < 0.02 | 0.06 | < 0.05 | 0.1 | 0.06 | 0.054261 | 0.03 | 0.05 | 0.04 | 0.05 | 0.05 | < 0.2 | 0.05 | 1.52 | < 0.2 | < 0.1 | < 5 | 0.007 | 0.04 | 0.04 | < 0.05 | [5 empty] | 196
| 1 | 6 | 0.03 | 0.06 | < 0.05 | 0.1 | 0.05 | 0.064494 | 0.03 | 0.05 | 0.06 | 0.05 | 0.04 | < 0.2 | < 0.05 | 1.05 | < 0.2 | < 0.1 | < 5 | 0.013 | 0.04 | 0.04 | 0.05
| Mean |  | 0.033333 | 0.058333 | 0.05 | 0.096667 | 0.0575 | 0.061305 | 0.033333 | 0.048333 | 0.04 | 0.05 | 0.046667 |  | 0.05 | 0.80167 |  |  |  | 0.039667 | 0.04 | 0.043333 | 0.056667
| Median |  | 0.03 | 0.06 | 0.05 | 0.1 | 0.06 | 0.064494 | 0.035 | 0.05 | 0.035 | 0.05 | 0.05 |  | 0.05 | 0.7 |  |  |  | 0.046 | 0.04 | 0.04 | 0.06
| Std Dev. |  | 0.0057735 | 0.01169 |  | 0.005164 | 0.005 | 0.0081685 | 0.008165 | 0.0040825 | 0.012649 | 7.6012e-18 | 0.005164 |  |  | 0.42372 |  |  |  | 0.024647 | 0 | 0.005164 | 0.0057735
| Rel.Std.Dev. |  | 17.32% | 20.04% |  | 5.342% | 8.696% | 13.32% | 24.49% | 8.447% | 31.62% | 1.52e-14% | 11.07% |  |  | 52.85% |  |  |  | 62.13% | 0% | 11.92% | 10.19%
| PDM3 |  | -30.07% | 22.38% | 4.894% | 102.8% | 20.63% | 28.61% | -30.07% | 1.398% | -16.08% | 4.894% | -2.099% |  | 4.894% | 1582% |  |  |  | -16.78% | -16.08% | -9.092% | 18.88%
| Z-Score (Absolute) |  | 1.96 | 0.92 | 1.37 | 3.93 | 0 | 0.67 | 1.05 | 0.13 | 0.52 | 0.26 | 0 | 4.19 | 1.37 | 59.28 | 4.19 | 0.26 | 192.64 | 0.55 | 0.52 | 0.26 | 0.46
| Analytical results for Th in OREAS 553b (Certified Value 6.18 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 2
| 1 | 1 | 6.85 | 6 | 7.7 | 7.69 | 6 | 5.6 | 6.7572 | 5.9 | 5.8 | 5.3 | 7.2 | 6.1 | 6.2727 | 5.9 | 5.8 | 6.2 | 6.1 | 5.467 | 3.53 | 8.54 | 5.53 | [5 empty] | 1
| 1 | 2 | 6.76 | 6.5 | 7.7 | 7.72 | 6.2 | 6 | 7.1834 | 5.9 | 5.9 | 5.4 | 6.6 | 5.8 | 6.7552 | 5.9 | 5.4 | 6.3 | 6.2 | 5.601 | 3.58 | 7.77 | 5.66 | [5 empty] | 10
| 1 | 3 | 6.63 | 6.3 | 6.6 | 7.62 | 6 | 5.8 | 6.9057 | 5.8 | 6 | 5.1 | 7 | 5.8 | 6.7941 | 5.9 | 5.5 | 6.2 | 6.3 | 5.416 | 3.68 | 8.4 | 5.63 | [5 empty] | 16
| 1 | 4 | 6.8 | 6 | 7.7 | 7.77 | 5.9 | 5.7 | 6.924 | 5.5 | 6 | 5.3 | 6.9 | 5.7 | 6.6682 | 6 | 5.4 | 6.2 | 6.6 | 5.497 | 3.66 | 8.44 | 5.46 | [5 empty] | 6.1773
| 1 | 5 | 6.78 | 5.6 | 6.4 | 7.48 | 5.8 | 5.9 | 6.3386 | 5.4 | 5.7 | 5.2 | 6.8 | 5.7 | 6.2489 | 5.9 | 5.4 | 6.2 | 6.6 | 5.385 | 3.58 | 8.27 | 5.57 | [5 empty] | 197
| 1 | 6 | 6.9 | 5.8 | 6.8 | 7.76 | 6 | 5.6 | 6.4514 | 5.8 | 5.9 | 5.3 | 7.3 | 5.9 | 6.2721 | 5.9 | 5.6 | 6.2 | 6.4 | 5.515 | 3.63 | 8.46 | 5.45
| Mean |  | 6.7867 | 6.0333 | 7.15 | 7.6733 | 5.9833 | 5.7667 | 6.7601 | 5.7167 | 5.8833 | 5.2667 | 6.9667 | 5.8333 | 6.5019 | 5.9167 | 5.5167 | 6.2167 | 6.3667 | 5.4802 | 3.61 | 8.3133 | 5.55
| Median |  | 6.79 | 6 | 7.25 | 7.705 | 6 | 5.75 | 6.8315 | 5.8 | 5.9 | 5.3 | 6.95 | 5.8 | 6.4704 | 5.9 | 5.45 | 6.2 | 6.35 | 5.482 | 3.605 | 8.42 | 5.55
| Std Dev. |  | 0.092014 | 0.3266 | 0.61563 | 0.10912 | 0.13292 | 0.1633 | 0.31631 | 0.2137 | 0.1169 | 0.10328 | 0.2582 | 0.15055 | 0.26327 | 0.040825 | 0.16021 | 0.040825 | 0.20656 | 0.07674 | 0.056569 | 0.28055 | 0.086487
| Rel.Std.Dev. |  | 1.356% | 5.413% | 8.61% | 1.422% | 2.221% | 2.832% | 4.679% | 3.738% | 1.987% | 1.961% | 3.706% | 2.581% | 4.049% | 0.69% | 2.904% | 0.6567% | 3.244% | 1.4% | 1.567% | 3.375% | 1.558%
| PDM3 |  | 9.865% | -2.331% | 15.75% | 24.22% | -3.14% | -6.648% | 9.434% | -7.457% | -4.759% | -14.74% | 12.78% | -5.568% | 5.254% | -4.219% | -10.69% | 0.6372% | 3.065% | -11.29% | -41.56% | 34.58% | -10.15%
| Z-Score (Absolute) |  | 1.16 | 0.07 | 1.69 | 2.44 | 0 | 0.31 | 1.12 | 0.39 | 0.14 | 1.04 | 1.42 | 0.22 | 0.75 | 0.1 | 0.67 | 0.34 | 0.55 | 0.73 | 3.43 | 3.37 | 0.63
| Analytical results for Ti in OREAS 553b (Indicative Value 0.001 wt.%) | [27 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  | 3
| 1 | 1 | 0.0008 | < 0.001 | < 0.01 | 0.0013 | < 0.005 | < 0.01 | 0.0010979 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.00083875 | < 0.01 | < 0.01 | 0.0018 | < 0.001 | < 0.005 | < 0.01 | 0.0015505 | 0.0024 | 0.0055 | 0.002 |  |  |  | 1
| 1 | 2 | 0.0008 | < 0.001 | < 0.01 | 0.0013 | < 0.005 | < 0.01 | 0.0013722 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.00089434 | < 0.01 | < 0.01 | 0.0021 | < 0.001 | < 0.005 | < 0.01 | 0.0016984 | < 0.0012 | 0.0054 | 0.002 |  |  |  | 26
| 1 | 3 | 0.0007 | < 0.001 | < 0.01 | 0.0013 | < 0.005 | < 0.01 | 0.0010773 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.000856 | < 0.01 | < 0.01 | 0.001 | < 0.001 | < 0.005 | < 0.01 | 0.0016408 | < 0.0012 | 0.006 | 0.002 |  |  |  | 16
| 1 | 4 | 0.0007 | < 0.001 | < 0.01 | 0.0014 | < 0.005 | < 0.01 | 0.001201 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.00089683 | < 0.01 | < 0.01 | 0.002 | < 0.001 | < 0.005 | < 0.01 | 0.0016128 | 0.003 | 0.0056 | 0.002 |  |  |  | 0.0014163
| 1 | 5 | 0.0007 | < 0.001 | < 0.01 | 0.0012 | < 0.005 | < 0.01 | 0.00099555 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.00084145 | < 0.01 | < 0.01 | 0.0015 | < 0.001 | < 0.005 | < 0.01 | 0.0015744 | 0.0012 | 0.0062 | 0.002 |  |  |  | 32
| 1 | 6 | 0.0008 | < 0.001 | < 0.01 | 0.0012 | < 0.005 | < 0.01 | 0.00099644 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.00084966 | < 0.01 | < 0.01 | 0.0021 | < 0.001 | < 0.005 | < 0.01 | 0.0015974 | 0.0012 | 0.0059 | 0.002
| Mean |  | 0.00075 |  |  | 0.0012833 |  |  | 0.0011234 | [5 empty] | 0.00086284 |  |  | 0.00175 |  |  |  | 0.0016124 | 0.00195 | 0.0057667 | 0.002
| Median |  | 0.00075 |  |  | 0.0013 |  |  | 0.0010876 | [5 empty] | 0.00085283 |  |  | 0.0019 |  |  |  | 0.0016051 | 0.0018 | 0.00575 | 0.002
| Std Dev. |  | 5.4772e-05 |  |  | 7.5277e-05 |  |  | 0.00014365 | [5 empty] | 2.6099e-05 |  |  | 0.00043243 |  |  |  | 5.2351e-05 | 0.0009 | 0.00031411 | 0
| Rel.Std.Dev. |  | 7.303% |  |  | 5.866% |  |  | 12.79% | [5 empty] | 3.025% |  |  | 24.71% |  |  |  | 3.247% | 46.15% | 5.447% | 0%
| PDM3 |  | -47.04% |  |  | -9.387% |  |  | -20.68% | [5 empty] | -39.08% |  |  | 23.56% |  |  |  | 13.85% | 37.68% | 307.2% | 41.21%
| Z-Score (Absolute) |  | 0.97 | 1.11 | 1.39 | 0.67 | 0 | 1.39 | 0.76 | 0 | 0 | 0 | 0 | 0 | 0.91 | 1.39 | 1.39 | 0.42 | 1.11 | 0 | 1.39 | 0.49 | 0.55 | 1.81 | 0.28
| Analytical results for Tl in OREAS 553b (Indicative Value 0.23 ppm) | [27 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.1g | 0.4g |  |  |  | 2
| 1 | 1 | 0.38 | 0.3 | < 0.02 | 0.41 | 0.21 | 0.16 | < 0.05 | 0.06 | 0.05 | 0.09 | 0.06 | 0.08 | 0.20924 | 0.22 | 0.14 | 0.5 | 0.361 | 0.3 | < 0.05 | 14.3 | 0.22 | 0.53 | 0.19 |  |  |  | 1
| 1 | 2 | 0.38 | 0.28 | < 0.02 | 0.43 | 0.2 | 0.14 | < 0.05 | 0.06 | 0.05 | 0.09 | 0.05 | 0.08 | 0.20874 | 0.22 | 0.14 | 0.52 | 0.351 | 0.3 | < 0.05 | 14.6 | 0.21 | 0.47 | 0.21 |  |  |  | 27
| 1 | 3 | 0.38 | 0.27 | < 0.02 | 0.43 | 0.2 | 0.15 | < 0.05 | 0.06 | 0.06 | 0.1 | 0.05 | 0.07 | 0.21081 | 0.22 | 0.14 | 0.49 | 0.356 | 0.3 | < 0.05 | 14.4 | 0.197 | 0.54 | 0.19 |  |  |  | 16
| 1 | 4 | 0.4 | 0.28 | < 0.02 | 0.44 | 0.21 | 0.16 | < 0.05 | 0.05 | 0.07 | 0.09 | 0.06 | 0.07 | 0.22486 | 0.22 | 0.14 | 0.5 | 0.369 | 0.3 | < 0.05 | 14.1 | 0.203 | 0.54 | 0.19 |  |  |  | 0.23012
| 1 | 5 | 0.39 | 0.27 | < 0.02 | 0.44 | 0.2 | 0.15 | < 0.05 | 0.06 | 0.07 | 0.08 | 0.05 | 0.07 | 0.20613 | 0.21 | 0.13 | 0.52 | 0.363 | 0.3 | < 0.05 | 14.3 | 0.219 | 0.5 | 0.19 |  |  |  | 33
| 1 | 6 | 0.4 | 0.26 | < 0.02 | 0.44 | 0.21 | 0.14 | < 0.05 | 0.06 | 0.05 | 0.11 | 0.05 | 0.07 | 0.20689 | 0.22 | 0.13 | 0.53 | 0.401 | 0.3 | < 0.05 | 14.9 | 0.224 | 0.52 | 0.18
| Mean |  | 0.38833 | 0.27667 |  | 0.43167 | 0.205 | 0.15 |  | 0.058333 | 0.058333 | 0.093333 | 0.053333 | 0.073333 | 0.21111 | 0.21833 | 0.13667 | 0.51 | 0.36683 | 0.3 |  | 14.4333 | 0.21217 | 0.51667 | 0.19167
| Median |  | 0.385 | 0.275 |  | 0.435 | 0.205 | 0.15 |  | 0.06 | 0.055 | 0.09 | 0.05 | 0.07 | 0.20899 | 0.22 | 0.14 | 0.51 | 0.362 | 0.3 |  | 14.35 | 0.2145 | 0.525 | 0.19
| Std Dev. |  | 0.0098319 | 0.013663 |  | 0.01169 | 0.0054772 | 0.0089443 |  | 0.0040825 | 0.0098319 | 0.010328 | 0.005164 | 0.005164 | 0.0069416 | 0.0040825 | 0.005164 | 0.015492 | 0.017826 | 0 |  | 0.28048 | 0.010647 | 0.027325 | 0.0098319
| Rel.Std.Dev. |  | 2.532% | 4.938% |  | 2.708% | 2.672% | 5.963% |  | 6.999% | 16.85% | 11.07% | 9.682% | 7.042% | 3.288% | 1.87% | 3.779% | 3.038% | 4.859% | 0% |  | 1.943% | 5.018% | 5.289% | 5.13%
| PDM3 |  | 68.75% | 20.23% |  | 87.58% | -10.92% | -34.82% |  | -74.65% | -74.65% | -59.44% | -76.82% | -68.13% | -8.262% | -5.123% | -40.61% | 121.6% | 59.41% | 30.37% |  | 6172% | -7.802% | 124.5% | -16.71%
| Z-Score (Absolute) |  | 0.92 | 0.38 | 0.91 | 1.12 | 0.03 | 0.23 | 0.83 | 0.67 | 0.67 | 0.51 | 0.7 | 0.6 | 0.06 | 0.1 | 0.3 | 1.5 | 0.81 | NA | 0.83 | 68.56 | 0.07 | 1.53 | 0.03
| Analytical results for Tm in OREAS 553b (Indicative Value 0.059 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 01 | 03 | 05 | 15 | 23 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [21 empty] | 3
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.25g | [21 empty] | 3
| 1 | 1 | 0.05 | < 0.1 | 0.09376 | 0.053121 | 0.04 | [21 empty] | 1
| 1 | 2 | 0.049 | < 0.1 | 0.09136 | 0.055057 | 0.044 | [21 empty] | 12
| 1 | 3 | 0.049 | < 0.1 | 0.09272 | 0.05894 | 0.04 | [21 empty] | 16
| 1 | 4 | 0.051 | < 0.1 | 0.09024 | 0.051052 | 0.04 | [21 empty] | 0.059369
| 1 | 5 | 0.051 | < 0.1 | 0.098 | 0.051679 | 0.042 | [21 empty] | 31
| 1 | 6 | 0.051 | < 0.1 | 0.08656 | 0.053364 | 0.042
| Mean |  | 0.050167 |  | 0.092107 | 0.053869 | 0.041333
| Median |  | 0.0505 |  | 0.09204 | 0.053242 | 0.041
| Std Dev. |  | 0.00098319 |  | 0.0038118 | 0.0028522 | 0.001633
| Rel.Std.Dev. |  | 1.96% |  | 4.138% | 5.295% | 3.951%
| PDM3 |  | -15.5% |  | 55.14% | -9.264% | -30.38%
| Z-Score (Absolute) |  | 0 | 0.03 | 7.64 | 0.67 | 1.61
| Analytical results for U in OREAS 553b (Certified Value 3.01 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 3
| 1 | 1 | 3.19 | 2.99 | 3.3 | 3.57 | 4.4178 | 2.85 | 2.99 | 3.3005 | 2.96 | 2.89 | 2.84 | 3.22 | 3.03 | 3.0694 | 2.74 | 2.89 | 3.87 | 3.1 | 2.9 | < 5 | 2.663 | 2.16 | 4.51 | 2.9 |  |  | 1
| 1 | 2 | 3.18 | 3.02 | 3.3 | 3.58 | 4.431 | 2.9 | 3.15 | 3.3313 | 2.97 | 2.89 | 2.86 | 3.03 | 2.94 | 3.2673 | 2.73 | 2.9 | 3.8 | 2.8 | 2.9 | < 5 | 2.767 | 2.2 | 4.23 | 2.95 |  |  | 13
| 1 | 3 | 3.11 | 2.89 | 3 | 3.57 | 4.468 | 2.95 | 3.1 | 3.3684 | 2.95 | 2.86 | 2.76 | 3.24 | 3.05 | 3.3314 | 2.8 | 2.96 | 3.79 | 2.9 | 2.9 | < 5 | 2.648 | 2.27 | 4.53 | 2.91 |  |  | 16
| 1 | 4 | 3.22 | 2.79 | 3.4 | 3.56 | 4.4965 | 2.88 | 3.04 | 3.3009 | 2.8 | 2.9 | 2.84 | 3.13 | 2.92 | 3.299 | 2.78 | 2.96 | 3.9 | 2.9 | 2.9 | < 5 | 2.625 | 2.24 | 4.66 | 2.88 |  |  | 3.0052
| 1 | 5 | 3.23 | 2.76 | 3 | 3.52 | 4.4641 | 2.84 | 2.98 | 3.4171 | 2.77 | 2.8 | 2.8 | 3.18 | 2.91 | 3.0638 | 2.75 | 2.94 | 3.8 | 2.8 | 2.8 | < 5 | 2.634 | 2.17 | 4.43 | 2.86 |  |  | 198
| 1 | 6 | 3.24 | 2.93 | 3.1 | 3.55 | 4.2974 | 2.91 | 3.18 | 3.3955 | 2.89 | 2.82 | 2.87 | 3.14 | 2.96 | 3.1227 | 2.73 | 2.97 | 3.91 | 3 | 2.9 | < 5 | 2.646 | 2.24 | 4.4 | 2.9
| Mean |  | 3.195 | 2.8967 | 3.1833 | 3.5583 | 4.4291 | 2.8883 | 3.0733 | 3.3523 | 2.89 | 2.86 | 2.8283 | 3.1567 | 2.9683 | 3.1923 | 2.755 | 2.9367 | 3.845 | 2.9167 | 2.8833 |  | 2.6638 | 2.2133 | 4.46 | 2.9
| Median |  | 3.205 | 2.91 | 3.2 | 3.565 | 4.4476 | 2.89 | 3.07 | 3.3498 | 2.92 | 2.875 | 2.84 | 3.16 | 2.95 | 3.195 | 2.745 | 2.95 | 3.835 | 2.9 | 2.9 |  | 2.647 | 2.22 | 4.47 | 2.9
| Std Dev. |  | 0.047645 | 0.10501 | 0.17224 | 0.02137 | 0.070374 | 0.040702 | 0.083347 | 0.049163 | 0.086487 | 0.041473 | 0.041191 | 0.075542 | 0.058452 | 0.12065 | 0.02881 | 0.033862 | 0.054681 | 0.1169 | 0.040825 |  | 0.052174 | 0.043665 | 0.14478 | 0.030332
| Rel.Std.Dev. |  | 1.491% | 3.625% | 5.411% | 0.6006% | 1.589% | 1.409% | 2.712% | 1.467% | 2.993% | 1.45% | 1.456% | 2.393% | 1.969% | 3.78% | 1.046% | 1.153% | 1.422% | 4.008% | 1.416% |  | 1.959% | 1.973% | 3.246% | 1.046%
| PDM3 |  | 6.317% | -3.611% | 5.928% | 18.41% | 47.38% | -3.888% | 2.268% | 11.55% | -3.833% | -4.831% | -5.885% | 5.041% | -1.226% | 6.225% | -8.325% | -2.28% | 27.95% | -2.945% | -4.054% |  | -11.36% | -26.35% | 48.41% | -3.5%
| Z-Score (Absolute) |  | 0.9 | 0.1 | 0.86 | 2.12 | 5.04 | 0.13 | 0.49 | 1.43 | 0.12 | 0.22 | 0.33 | 0.77 | 0.14 | 0.89 | 0.58 | 0.03 | 3.08 | 0.03 | 0.15 | 1.43 | 0.88 | 2.4 | 5.15 | 0.09
| Analytical results for V in OREAS 553b (Certified Value 56 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  | 0
| 1 | 1 | 69 | 54 | 88 | 88 | 62 | 61 | 47 | 57.7193 | 55 | 52 | 51 | 60 | 50 | 59.9117 | 45.2 | 61 | 115 | 56 | 65 | 43.5 | 56.18 | 34 | 136 | 50 |  |  | 1
| 1 | 2 | 69 | 60 | 90 | 89 | 65 | 62 | 48 | 64.0537 | 54 | 53 | 51 | 58 | 50 | 63.6553 | 44 | 61 | 121 | 54 | 64 | 43.6 | 57.47 | 34.6 | 113 | 51 |  |  | 29
| 1 | 3 | 70 | 60 | 76 | 86 | 64 | 61 | 50 | 59.6487 | 55 | 53 | 50 | 59 | 52 | 63.4609 | 46.7 | 61 | 116 | 54 | 64 | 44.2 | 55.775 | 36.2 | 142 | 50 |  |  | 16
| 1 | 4 | 70 | 55 | 86 | 87 | 66 | 62 | 49 | 65.0617 | 52 | 52 | 51 | 60 | 50 | 61.6874 | 45.2 | 63 | 119 | 51 | 63 | 44 | 55.867 | 35.9 | 135 | 48 |  |  | 55.7682
| 1 | 5 | 67 | 51 | 70 | 83 | 64 | 62 | 47 | 57.6771 | 53 | 52 | 52 | 59 | 49 | 59.9029 | 44.8 | 61 | 121 | 51 | 63 | 44.2 | 56.688 | 34.1 | 127 | 48 |  |  | 199
| 1 | 6 | 71 | 54 | 75 | 84 | 65 | 61 | 50 | 60.0941 | 55 | 50 | 52 | 61 | 50 | 58.6748 | 44.8 | 62 | 120 | 55 | 64 | 43.7 | 56.762 | 36.8 | 121 | 52
| Mean |  | 69.3333 | 55.6667 | 80.8333 | 86.1667 | 64.3333 | 61.5 | 48.5 | 60.7091 | 54 | 52 | 51.1667 | 59.5 | 50.1667 | 61.2155 | 45.1167 | 61.5 | 118.6667 | 53.5 | 63.8333 | 43.8667 | 56.457 | 35.2667 | 129 | 49.8333
| Median |  | 69.5 | 54.5 | 81 | 86.5 | 64.5 | 61.5 | 48.5 | 59.8714 | 54.5 | 52 | 51 | 59.5 | 50 | 60.7995 | 45 | 61 | 119.5 | 54 | 64 | 43.85 | 56.434 | 35.25 | 131 | 50
| Std Dev. |  | 1.3663 | 3.6148 | 8.2077 | 2.3166 | 1.3663 | 0.54772 | 1.3784 | 3.1549 | 1.2649 | 1.0954 | 0.75277 | 1.0488 | 0.98319 | 2.054 | 0.89088 | 0.83666 | 2.582 | 2.0736 | 0.75277 | 0.30768 | 0.64226 | 1.186 | 10.7517 | 1.6021
| Rel.Std.Dev. |  | 1.971% | 6.494% | 10.15% | 2.689% | 2.124% | 0.8906% | 2.842% | 5.197% | 2.342% | 2.107% | 1.471% | 1.763% | 1.96% | 3.355% | 1.975% | 1.36% | 2.176% | 3.876% | 1.179% | 0.7014% | 1.138% | 3.363% | 8.335% | 3.215%
| PDM3 |  | 24.32% | -0.182% | 44.95% | 54.51% | 15.36% | 10.28% | -13.03% | 8.86% | -3.171% | -6.757% | -8.251% | 6.692% | -10.04% | 9.768% | -19.1% | 10.28% | 112.8% | -4.067% | 14.46% | -21.34% | 1.235% | -36.76% | 131.3% | -10.64%
| Z-Score (Absolute) |  | 1.16 | 0.24 | 2.34 | 2.89 | 0.65 | 0.36 | 0.97 | 0.28 | 0.41 | 0.61 | 0.7 | 0.16 | 0.8 | 0.33 | 1.32 | 0.36 | 6.22 | 0.46 | 0.6 | 1.45 | 0.16 | 2.33 | 7.27 | 0.83
| Analytical results for W in OREAS 553b (Certified Value 0.96 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  |  |  | 3
| 1 | 1 | 0.9 | 0.82 | 0.8 | 0.6 | < 10 | 1.38 | 1.1 | 0.90464 | 1.07 | 0.98 | 1.12 | 0.86 | 0.99 | 0.95047 | 0.8 | 0.7 | 0.7 | 0.6 | < 5 | 0.957 | 2 | 1.04 | 1.38 |  |  |  | 1
| 1 | 2 | 0.88 | 0.73 | 0.9 | 0.6 | < 10 | 1.39 | 1 | 0.99659 | 1.01 | 0.97 | 1.03 | 0.88 | 1.08 | 1.0049 | 0.8 | 0.7 | 0.6 | 0.6 | < 5 | 0.955 | 2 | 0.95 | 1.47 |  |  |  | 14
| 1 | 3 | 0.87 | 0.71 | 0.8 | 0.6 | < 10 | 1.32 | 1.1 | 0.90339 | 1 | 1 | 0.97 | 0.86 | 0.98 | 1.0932 | 0.9 | 0.5 | 0.6 | 0.6 | < 5 | 0.961 | 2 | 1.04 | 1.31 |  |  |  | 16
| 1 | 4 | 0.92 | 0.74 | 0.8 | 0.6 | < 10 | 1.32 | 1 | 0.97172 | 0.93 | 0.97 | 1.06 | 0.87 | 0.96 | 1.0122 | 0.8 | 0.7 | 0.7 | 0.6 | < 5 | 1.008 | 2.1 | 1.01 | 1.36 |  |  |  | 0.96473
| 1 | 5 | 0.86 | 0.78 | 0.8 | 0.6 | < 10 | 1.37 | 1.1 | 0.85952 | 0.97 | 0.96 | 1.01 | 0.85 | 0.94 | < 1 | 0.8 | 0.6 | 0.7 | 0.6 | < 5 | 0.98 | 1.9 | 1.02 | 1.39 |  |  |  | 200
| 1 | 6 | 0.9 | 0.75 | 0.8 | 0.5 | < 10 | 1.4 | 1.1 | 0.85922 | 0.95 | 0.94 | 0.99 | 0.84 | 0.95 | 1.0148 | 0.9 | 0.7 | 0.7 | 0.6 | < 5 | 1.007 | 2 | 1.05 | 1.29
| Mean |  | 0.88833 | 0.755 | 0.81667 | 0.58333 |  | 1.3633 | 1.0667 | 0.91585 | 0.98833 | 0.97 | 1.03 | 0.86 | 0.98333 | 1.0151 | 0.83333 | 0.65 | 0.66667 | 0.6 |  | 0.978 | 2 | 1.0183 | 1.3667
| Median |  | 0.89 | 0.745 | 0.8 | 0.6 |  | 1.375 | 1.1 | 0.90402 | 0.985 | 0.97 | 1.02 | 0.86 | 0.97 | 1.0122 | 0.8 | 0.7 | 0.7 | 0.6 |  | 0.9705 | 2 | 1.03 | 1.37
| Std Dev. |  | 0.022286 | 0.03937 | 0.040825 | 0.040825 |  | 0.035024 | 0.05164 | 0.0571 | 0.049967 | 0.02 | 0.054037 | 0.014142 | 0.050859 | 0.050967 | 0.05164 | 0.083666 | 0.05164 | 0 |  | 0.024511 | 0.063246 | 0.03656 | 0.064083
| Rel.Std.Dev. |  | 2.509% | 5.215% | 4.999% | 6.999% |  | 2.569% | 4.841% | 6.235% | 5.056% | 2.062% | 5.246% | 1.644% | 5.172% | 5.021% | 6.197% | 12.87% | 7.746% | 0% |  | 2.506% | 3.162% | 3.59% | 4.689%
| PDM3 |  | -7.919% | -21.74% | -15.35% | -39.53% |  | 41.32% | 10.57% | -5.067% | 2.447% | 0.5463% | 6.766% | -10.86% | 1.928% | 5.223% | -13.62% | -32.62% | -30.9% | -37.81% |  | 1.376% | 107.3% | 5.556% | 41.66%
| Z-Score (Absolute) |  | 0.95 | 2.28 | NA | NA | 40.13 | 3.8 | NA | 0.67 | 0.05 | 0.13 | 0.47 | 1.23 | 0 | 0.54 | NA | NA | NA | NA | 15.15 | 0.05 | NA | 0.35 | 3.83
| Analytical results for Y in OREAS 553b (Certified Value 3.1 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 2
| 1 | 1 | 3.41 | 3.89 | 4.67 | 8.9 | 3.31 | 3.06 | 3.4519 | 3.17 | 3.02 | 3 | 3.5 | 3.19 | 3.4859 | 2.8 | 3.13 | 2.9 | < 5 | 2.999 | 2.37 | 6.23 | 2.37 | [5 empty] | 1
| 1 | 2 | 3.43 | 3.98 | 4.61 | 8.1 | 3.47 | 2.87 | 3.714 | 3.06 | 3.09 | 3.05 | 3.39 | 3.18 | 3.5694 | 2.8 | 3.26 | 2.9 | < 5 | 3.057 | 2.37 | 6.26 | 2.35 | [5 empty] | 15
| 1 | 3 | 3.44 | 3.54 | 4.57 | 9.5 | 3.35 | 2.91 | 3.5369 | 3.12 | 3.07 | 2.84 | 3.43 | 3.2 | 3.505 | 2.8 | 3.23 | 2.9 | < 5 | 2.946 | 2.43 | 6.52 | 2.23 | [5 empty] | 16
| 1 | 4 | 3.43 | 3.69 | 4.63 | 9.2 | 3.39 | 3.03 | 3.6819 | 2.84 | 3.02 | 2.99 | 3.49 | 3.16 | 3.43 | 2.8 | 3.32 | 2.8 | < 5 | 2.99 | 2.41 | 5.98 | 2.33 | [5 empty] | 3.0963
| 1 | 5 | 3.37 | 3.37 | 4.64 | 9.7 | 3.38 | 2.82 | 3.4126 | 2.96 | 2.9 | 2.96 | 3.38 | 2.98 | 3.416 | 2.8 | 3.22 | 2.8 | < 5 | 2.993 | 2.37 | 5.71 | 2.36 | [5 empty] | 201
| 1 | 6 | 3.5 | 3.52 | 4.67 | 9.2 | 3.31 | 3.08 | 3.5659 | 3 | 2.91 | 2.97 | 3.39 | 3 | 3.3864 | 2.8 | 3.33 | 2.9 | < 5 | 2.975 | 2.42 | 5.44 | 2.4
| Mean |  | 3.43 | 3.665 | 4.6317 | 9.1 | 3.3683 | 2.9617 | 3.5605 | 3.025 | 3.0017 | 2.9683 | 3.43 | 3.1183 | 3.4655 | 2.8 | 3.2483 | 2.8667 |  | 2.9933 | 2.395 | 6.0233 | 2.34
| Median |  | 3.43 | 3.615 | 4.635 | 9.2 | 3.365 | 2.97 | 3.5514 | 3.03 | 3.02 | 2.98 | 3.41 | 3.17 | 3.458 | 2.8 | 3.245 | 2.9 |  | 2.9915 | 2.39 | 6.105 | 2.355
| Std Dev. |  | 0.042426 | 0.23416 | 0.038166 | 0.56214 | 0.060139 | 0.10907 | 0.12048 | 0.11862 | 0.079854 | 0.070261 | 0.053292 | 0.10048 | 0.067442 | 4.8648e-16 | 0.073598 | 0.05164 |  | 0.036533 | 0.028107 | 0.39642 | 0.058652
| Rel.Std.Dev. |  | 1.237% | 6.389% | 0.824% | 6.177% | 1.785% | 3.683% | 3.384% | 3.921% | 2.66% | 2.367% | 1.554% | 3.222% | 1.946% | 1.737e-14% | 2.266% | 1.801% |  | 1.22% | 1.174% | 6.581% | 2.506%
| PDM3 |  | 10.78% | 18.37% | 49.59% | 193.9% | 8.785% | -4.349% | 14.99% | -2.304% | -3.057% | -4.134% | 10.78% | 0.7105% | 11.92% | -9.57% | 4.909% | -7.417% |  | -3.327% | -22.65% | 94.53% | -24.43%
| Z-Score (Absolute) |  | 0.67 | 1.18 | 3.27 | 12.94 | 0.54 | 0.34 | 0.96 | 0.2 | 0.25 | 0.32 | 0.67 | 0 | 0.75 | 0.69 | 0.28 | 0.54 | 1.34 | 0.27 | 1.56 | 6.29 | 1.68
| Analytical results for Yb in OREAS 553b (Certified Value 0.28 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 05 | 08 | 15 | 16 | 17 | 22 | 23 | 24 | 26 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | [15 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.4g | [15 empty] | 2
| 1 | 1 | 0.337 | 0.3 | 0.59656 | 0.3 | 0.31335 | 0.26 | 0.3 | < 5 | 0.275 | 0.24 | 0.3 | [15 empty] | 1
| 1 | 2 | 0.319 | 0.4 | 0.56712 | 0.3 | 0.34371 | 0.22 | 0.3 | < 5 | 0.265 | 0.24 | 0.3 | [15 empty] | 16
| 1 | 3 | 0.304 | 0.3 | 0.58712 | 0.3 | 0.35959 | 0.24 | 0.3 | < 5 | 0.266 | 0.25 | 0.3 | [15 empty] | 16
| 1 | 4 | 0.33 | 0.4 | 0.59048 | 0.3 | 0.31861 | 0.25 | 0.3 | < 5 | 0.281 | 0.25 | 0.3 | [15 empty] | 0.28194
| 1 | 5 | 0.297 | 0.3 | 0.5704 | 0.3 | 0.33291 | 0.27 | 0.3 | < 5 | 0.271 | 0.23 | 0.3 | [15 empty] | 202
| 1 | 6 | 0.322 | 0.3 | 0.56096 | 0.3 | 0.30993 | 0.25 | 0.3 | < 5 | 0.263 | 0.25 | 0.3
| Mean |  | 0.31817 | 0.33333 | 0.57877 | 0.3 | 0.32968 | 0.24833 | 0.3 |  | 0.27017 | 0.24333 | 0.3
| Median |  | 0.3205 | 0.3 | 0.57876 | 0.3 | 0.32576 | 0.25 | 0.3 |  | 0.2685 | 0.245 | 0.3
| Std Dev. |  | 0.015224 | 0.05164 | 0.014466 | 0 | 0.019398 | 0.017224 | 0 |  | 0.0068823 | 0.008165 | 0
| Rel.Std.Dev. |  | 4.785% | 15.49% | 2.499% | 0% | 5.884% | 6.936% | 0% |  | 2.547% | 3.355% | 0%
| PDM3 |  | 12.85% | 18.23% | 105.3% | 6.407% | 16.94% | -11.92% | 6.407% |  | -4.175% | -13.69% | 6.407%
| Z-Score (Absolute) |  | 0 | NA | 2.52 | NA | 0.11 | 0.67 | NA | 21.07 | 0.46 | 0.72 | NA
Results from laboratories 3, 8, 17 and 26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 553b (Certified Value 13 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g |  | 2
| 1 | 1 | 14 | 20.3 | 12.4 | 12 | 10 | 13 | 12 | 12.4348 | 13 | 12 | 13 | 13 | 13 | 13.5403 | 12.7 | 13 | 14 | 18 | 13 | 17 | 15.5 | 16.53 | 12.8 | 12.8 | 13 |  | 1
| 1 | 2 | 13 | 20.5 | 13.3 | 12 | 10 | 15 | 14 | 12.4848 | 12 | 12 | 13 | 13 | 12 | 14.0721 | 12.7 | 12 | 15 | 18 | 12 | 12 | 18.7 | 18.873 | 14.4 | 13.4 | 14 |  | 17
| 1 | 3 | 13 | 19.9 | 13.7 | 12 | 10 | 15 | 13 | 11.7112 | 13 | 12 | 13 | 13 | 12 | 13.2793 | 14 | 13 | 15 | 19 | 12 | 13 | 15.4 | 20.783 | 14.9 | 12.7 | 14 |  | 16
| 1 | 4 | 13 | 19.3 | 13.5 | 11 | 11 | 14 | 14 | 11.466 | 13 | 11 | 12 | 13 | 12 | 13.3343 | 12.9 | 13 | 15 | 18 | 12 | 13 | 15.3 | 18.701 | 15.1 | 12.2 | 15 |  | 13.0011
| 1 | 5 | 13 | 19.3 | 12.5 | 13 | 11 | 14 | 12 | 11.2949 | 13 | 12 | 17 | 13 | 12 | 12.9459 | 13.3 | 12 | 15 | 17 | 12 | 12 | 14.3 | 18.596 | 14.5 | 12.3 | 15 |  | 203
| 1 | 6 | 13 | 18.2 | 13.1 | 13 | 10 | 14 | 13 | 11.6619 | 13 | 11 | 14 | 13 | 12 | 13.1235 | 13 | 13 | 15 | 17 | 12 | 13 | 15.5 | 18.131 | 15.3 | 12.5 | 16
| Mean |  | 13.1667 | 19.5833 | 13.0833 | 12.1667 | 10.3333 | 14.1667 | 13 | 11.8422 | 12.8333 | 11.6667 | 13.6667 | 13 | 12.1667 | 13.3826 | 13.1 | 12.6667 | 14.8333 | 17.8333 | 12.1667 | 13.3333 | 15.7833 | 18.6023 | 14.5 | 12.65 | 14.5
| Median |  | 13 | 19.6 | 13.2 | 12 | 10 | 14 | 13 | 11.6865 | 13 | 12 | 13 | 13 | 12 | 13.3068 | 12.95 | 13 | 15 | 18 | 12 | 13 | 15.45 | 18.6485 | 14.7 | 12.6 | 14.5
| Std Dev. |  | 0.40825 | 0.84004 | 0.53072 | 0.75277 | 0.5164 | 0.75277 | 0.89443 | 0.50097 | 0.40825 | 0.5164 | 1.7512 | 0 | 0.40825 | 0.39273 | 0.49396 | 0.5164 | 0.40825 | 0.75277 | 0.40825 | 1.8619 | 1.4999 | 1.3678 | 0.90111 | 0.43243 | 1.0488
| Rel.Std.Dev. |  | 3.101% | 4.29% | 4.056% | 6.187% | 4.997% | 5.314% | 6.88% | 4.23% | 3.181% | 4.426% | 12.81% | 0% | 3.355% | 2.935% | 3.771% | 4.077% | 2.752% | 4.221% | 3.355% | 13.96% | 9.503% | 7.353% | 6.215% | 3.418% | 7.233%
| PDM3 |  | 1.273% | 50.63% | 0.6323% | -6.418% | -20.52% | 8.965% | -0.008678% | -8.914% | -1.291% | -10.26% | 5.119% | -0.008678% | -6.418% | 2.934% | 0.7605% | -2.573% | 14.09% | 37.17% | -6.418% | 2.555% | 21.4% | 43.08% | 11.53% | -2.701% | 11.53%
| Z-Score (Absolute) |  | 0.05 | 4.68 | 0.01 | 0.67 | 2 | 0.77 | 0.07 | 0.91 | 0.19 | 1.04 | 0.41 | 0.07 | 0.67 | 0.2 | 0 | 0.31 | 1.25 | 3.42 | 0.67 | 0.17 | 1.94 | 3.98 | 1.01 | 0.33 | 1.01
| Analytical results for Zr in OREAS 553b (Certified Value 9.39 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 22 | 23 | 24 | 25 | 26 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.1g | 0.4g | [5 empty] | 2
| 1 | 1 | 10 | 11.7 | 9.2 | 10.1 | 11.4 | 8.8 | 8.8081 | 8.5 | 8.7 | 8.5 | 9.9 | 8.2 | 9.5618 | 8.4 | 9.3 | 8 | 10.8 | 9.303 | 6.8 | 12.8 | 7.7 | [5 empty] | 1
| 1 | 2 | 9.7 | 11.5 | 9.5 | 10.4 | 11.1 | 9 | 11.5846 | 8.3 | 9.1 | 8.6 | 8.8 | 8.4 | 10.0117 | 8.1 | 9.3 | 7.9 | 10.3 | 9.601 | 7.09 | 13.9 | 8 | [5 empty] | 21
| 1 | 3 | 9.8 | 10.5 | 9.4 | 9.9 | 10.9 | 8.5 | 9.3273 | 8.3 | 9.3 | 8.2 | 9.2 | 8.3 | 9.8976 | 8.4 | 9.4 | 8 | 10.4 | 9.272 | 6.98 | 13.5 | 8.1 | [5 empty] | 16
| 1 | 4 | 9.9 | 11.6 | 9.6 | 10.7 | 10.9 | 9.2 | 12.3528 | 7.6 | 8.8 | 8.4 | 9.6 | 8.3 | 9.4574 | 8.3 | 9.6 | 8.1 | 10.4 | 9.535 | 7.08 | 13.4 | 7.8 | [5 empty] | 9.3918
| 1 | 5 | 9.8 | 10.3 | 9.5 | 10.6 | 10.9 | 8.6 | 9.8279 | 7.9 | 8.3 | 8.3 | 9.4 | 7.7 | 9.561 | 8.5 | 9.3 | 8.3 | 10.8 | 9.479 | 6.89 | 12.4 | 7.9 | [5 empty] | 204
| 1 | 6 | 9.9 | 10.9 | 9.7 | 10.5 | 10.6 | 9 | 11.174 | 7.9 | 8.4 | 8.4 | 9.5 | 7.8 | 9.617 | 8.3 | 9.5 | 8.3 | 10.6 | 9.544 | 7.05 | 12.2 | 8.4
| Mean |  | 9.85 | 11.0833 | 9.4833 | 10.3667 | 10.9667 | 8.85 | 10.5124 | 8.0833 | 8.7667 | 8.4 | 9.4 | 8.1167 | 9.6844 | 8.3333 | 9.4 | 8.1 | 10.55 | 9.4557 | 6.9817 | 13.0333 | 7.9833
| Median |  | 9.85 | 11.2 | 9.5 | 10.45 | 10.9 | 8.9 | 10.5009 | 8.1 | 8.75 | 8.4 | 9.45 | 8.25 | 9.5894 | 8.35 | 9.35 | 8.05 | 10.5 | 9.507 | 7.015 | 13.1 | 7.95
| Std Dev. |  | 0.10488 | 0.60139 | 0.17224 | 0.30768 | 0.26583 | 0.26646 | 1.3966 | 0.33714 | 0.38816 | 0.14142 | 0.37417 | 0.29269 | 0.2186 | 0.13663 | 0.12649 | 0.16733 | 0.21679 | 0.13624 | 0.11618 | 0.67132 | 0.24833
| Rel.Std.Dev. |  | 1.065% | 5.426% | 1.816% | 2.968% | 2.424% | 3.011% | 13.28% | 4.171% | 4.428% | 1.684% | 3.98% | 3.606% | 2.257% | 1.64% | 1.346% | 2.066% | 2.055% | 1.441% | 1.664% | 5.151% | 3.111%
| PDM3 |  | 4.879% | 18.01% | 0.9748% | 10.38% | 16.77% | -5.769% | 11.93% | -13.93% | -6.656% | -10.56% | 0.08745% | -13.58% | 3.116% | -11.27% | 0.08745% | -13.75% | 12.33% | 0.6802% | -25.66% | 38.77% | -15%
| Z-Score (Absolute) |  | 0.28 | 1.06 | 0.05 | 0.61 | 0.99 | 0.35 | 0.7 | 0.83 | 0.4 | 0.63 | 0 | 0.81 | 0.18 | 0.67 | 0 | 0.82 | 0.73 | 0.04 | 1.53 | 2.3 | 0.9
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C722: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C758: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C776: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C812: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C830: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C848: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C866: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C884: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C902: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C920: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C938: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C956: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C975: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C993: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1011: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1029: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1047: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1065: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1083: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1119: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1137: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1156: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1174: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for As in OREAS 553b (Indicative Value < 10 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 5
| 1 | 3 | < 10 | 16
| 1 | 4 | < 10 | < 10
| 1 | 5 | < 10 | 35
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 553b (Indicative Value 0.564 wt.%) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 24 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.571 | 1
| 1 | 2 | 0.56 | 30
| 1 | 3 | 0.57 | 16
| 1 | 4 | 0.543 | 0.56417
| 1 | 5 | 0.565 | 36
| 1 | 6 | 0.576
| Mean |  | 0.56417
| Median |  | 0.5675
| Std Dev. |  | 0.01172
| Rel.Std.Dev. |  | 2.077%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 553b (Indicative Value 3.6 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 24 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.57 | 1
| 1 | 2 | 3.57 | 31
| 1 | 3 | 3.6 | 16
| 1 | 4 | 3.58 | 3.5967
| 1 | 5 | 3.63 | 37
| 1 | 6 | 3.63
| Mean |  | 3.5967
| Median |  | 3.59
| Std Dev. |  | 0.028048
| Rel.Std.Dev. |  | 0.7798%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 553b (Indicative Value 102 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 50 | 29
| 1 | 3 | 80 | 16
| 1 | 4 | 270 | 101.6667
| 1 | 5 | 110 | 35
| 1 | 6 | 40
| Mean |  | 101.6667
| Median |  | 70
| Std Dev. |  | 86.1201
| Rel.Std.Dev. |  | 84.71%
| PDM3 |  | -3.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 553b (Indicative Value 157 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 180 | 13
| 1 | 3 | 80 | 16
| 1 | 4 | 170 | 156.6667
| 1 | 5 | 280 | 36
| 1 | 6 | 140
| Mean |  | 156.6667
| Median |  | 155
| Std Dev. |  | 72.8469
| Rel.Std.Dev. |  | 46.5%
| PDM3 |  | -2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 553b (Indicative Value 3.32 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 24 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.33 | 1
| 1 | 2 | 3.3 | 31
| 1 | 3 | 3.34 | 16
| 1 | 4 | 3.3 | 3.32
| 1 | 5 | 3.28 | 37
| 1 | 6 | 3.37
| Mean |  | 3.32
| Median |  | 3.315
| Std Dev. |  | 0.032863
| Rel.Std.Dev. |  | 0.9899%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 553b (Indicative Value 102 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 70 | 29
| 1 | 3 | 100 | 16
| 1 | 4 | 120 | 101.6667
| 1 | 5 | 110 | 35
| 1 | 6 | 130
| Mean |  | 101.6667
| Median |  | 105
| Std Dev. |  | 23.1661
| Rel.Std.Dev. |  | 22.79%
| PDM3 |  | -3.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 553b (Indicative Value 103 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 24 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 110 | 1
| 1 | 2 | 90 | 17
| 1 | 3 | 100 | 16
| 1 | 4 | 100 | 103.3333
| 1 | 5 | 100 | 36
| 1 | 6 | 120
| Mean |  | 103.3333
| Median |  | 100
| Std Dev. |  | 10.328
| Rel.Std.Dev. |  | 9.995%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.