## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 313
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 75.573 | 73.8969 | 77.2491 | 74.3189 | 76.8271
Al, Aluminium (wt.%) | 4.8992 | 4.7784 | 5.0199 | 4.832 | 4.9663
As, Arsenic (ppm) | 93.8022 | 90.6771 | 96.9273 | 91.2173 | 96.3871
Be, Beryllium (ppm) | 2.1866 | 2.0682 | 2.305 | 2.1155 | 2.2576
Bi, Bismuth (ppm) | 0.43355 | 0.40758 | 0.45952 | 0.41484 | 0.45226
Ca, Calcium (wt.%) | 1.3547 | 1.3257 | 1.3837 | 1.3287 | 1.3807
Cd, Cadmium (ppm) | 113.1288 | 109.1642 | 117.0933 | 110.9455 | 115.312
Ce, Cerium (ppm) | 56.1653 | 52.9602 | 59.3705 | 54.6166 | 57.7141
Co, Cobalt (ppm) | 25.5278 | 24.5682 | 26.4874 | 24.9785 | 26.0771
Cr, Chromium (ppm) | 68.9483 | 63.8323 | 74.0643 | 66.3061 | 71.5905
Cs, Caesium (ppm) | 3.1218 | 2.9966 | 3.2471 | 3.0228 | 3.2209
Cu, Copper (wt.%) | 0.035683 | 0.034979 | 0.036386 | 0.035116 | 0.036249
Dy, Dysprosium (ppm) | 2.6886 | 2.252 | 3.1252 | 2.5583 | 2.8189
Er, Erbium (ppm) | 1.3424 | 1.1231 | 1.5618 | 1.2776 | 1.4073
Eu, Europium (ppm) | 0.90384 | 0.76913 | 1.0385 | 0.86336 | 0.94432
Fe, Iron (wt.%) | 6.1545 | 6.0279 | 6.2812 | 6.0573 | 6.2517
Ga, Gallium (ppm) | 14.2576 | 13.6174 | 14.8978 | 13.8288 | 14.6864
Gd, Gadolinium (ppm) | 3.8569 | 3.2646 | 4.4491 | 3.705 | 4.0087
Ge, Germanium (ppm) | 0.13559 | 0.10563 | 0.16555 | IND | IND
Hf, Hafnium (ppm) | 2.0754 | 1.9617 | 2.1891 | 1.9814 | 2.1695
Ho, Holmium (ppm) | 0.49743 | 0.45276 | 0.5421 | 0.4661 | 0.52876
In, Indium (ppm) | 0.29068 | 0.27195 | 0.30941 | 0.2784 | 0.30296
K, Potassium (wt.%) | 1.9987 | 1.9616 | 2.0358 | 1.9616 | 2.0358
La, Lanthanum (ppm) | 28.8871 | 26.9323 | 30.8418 | 27.8709 | 29.9032
Li, Lithium (ppm) | 23.4667 | 22.4646 | 24.4687 | 22.9958 | 23.9375
Lu, Lutetium (ppm) | 0.21098 | 0.16632 | 0.25564 | 0.19189 | 0.23007
Mg, Magnesium (wt.%) | 1.0263 | 1.0049 | 1.0476 | 1.0064 | 1.0461
Mn, Manganese (wt.%) | 0.18567 | 0.18117 | 0.19016 | 0.18231 | 0.18902
Mo, Molybdenum (ppm) | 5.2571 | 5.0211 | 5.4931 | 5.0738 | 5.4404
Na, Sodium (wt.%) | 0.075703 | 0.071045 | 0.080362 | 0.073903 | 0.077503
Nb, Niobium (ppm) | 3.3011 | 2.9628 | 3.6395 | 3.0992 | 3.5031
Nd, Neodymium (ppm) | 26.5575 | 24.3007 | 28.8143 | 25.9198 | 27.1952
Ni, Nickel (ppm) | 15.9264 | 15.0398 | 16.813 | 15.3126 | 16.5401
P, Phosphorus (wt.%) | 0.031266 | 0.030251 | 0.032282 | 0.030353 | 0.03218
Pb, Lead (wt.%) | 3.8963 | 3.804 | 3.9885 | 3.8378 | 3.9548
Pr, Praseodymium (ppm) | 6.831 | 6.0662 | 7.5958 | 6.632 | 7.03
Rb, Rubidium (ppm) | 107.6525 | 103.5969 | 111.7081 | 105.1853 | 110.1198
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 4.4423 | 4.3141 | 4.5704 | 4.3745 | 4.51
Sb, Antimony (ppm) | 44.4058 | 42.3673 | 46.4442 | 42.9759 | 45.8356
Sc, Scandium (ppm) | 8.7012 | 8.2882 | 9.1143 | 8.4763 | 8.9261
Sm, Samarium (ppm) | 4.8749 | 4.5347 | 5.2152 | 4.6828 | 5.067
Sn, Tin (ppm) | 2.5852 | 2.4251 | 2.7453 | 2.4427 | 2.7278
Sr, Strontium (ppm) | 41.676 | 40.175 | 43.1771 | 40.334 | 43.0181
Ta, Tantalum (ppm) | 0.2391 | 0.19476 | 0.28345 | 0.21624 | 0.26197
Tb, Terbium (ppm) | 0.50888 | 0.43626 | 0.58151 | 0.4755 | 0.54227
Te, Tellurium (ppm) | 0.084432 | 0.06474 | 0.10412 | 0.073096 | 0.095767
Th, Thorium (ppm) | 9.9832 | 9.4725 | 10.4938 | 9.6455 | 10.3208
Ti, Titanium (wt.%) | 0.15459 | 0.14766 | 0.16151 | 0.1507 | 0.15848
Tl, Thallium (ppm) | 20.0787 | 19.2664 | 20.8909 | 19.5871 | 20.5702
Tm, Thulium (ppm) | 0.17925 | 0.1314 | 0.2271 | IND | IND
U, Uranium (ppm) | 3.4476 | 3.275 | 3.6201 | 3.3447 | 3.5504
V, Vanadium (ppm) | 133.2304 | 129.7975 | 136.6632 | 130.1137 | 136.3471
W, Tungsten (ppm) | 1.3698 | 1.2328 | 1.5068 | 1.2237 | 1.5159
Y, Yttrium (ppm) | 13.499 | 12.431 | 14.5669 | 13.0498 | 13.9482
Yb, Ytterbium (ppm) | 1.4549 | 1.2434 | 1.6663 | 1.3827 | 1.527
Zn, Zinc (wt.%) | 3.5482 | 3.4699 | 3.6264 | 3.4987 | 3.5976
Zr, Zirconium (ppm) | 69.8741 | 66.9335 | 72.8148 | 68.0703 | 71.678
Aqua Regia Digestion
Ag, Silver (ppm) | 74.5965 | 72.8407 | 76.3522 | 73.3521 | 75.8408
Al, Aluminium (wt.%) | 0.78732 | 0.74456 | 0.83007 | 0.7681 | 0.80654
As, Arsenic (ppm) | 93.3309 | 87.9422 | 98.7196 | 91.3631 | 95.2987
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.5098 | 0.47534 | 0.54426 | 0.48986 | 0.52974
Bi, Bismuth (ppm) | 0.31715 | 0.29136 | 0.34295 | 0.29699 | 0.33732
Ca, Calcium (wt.%) | 1.3414 | 1.309 | 1.3737 | 1.3075 | 1.3753
Cd, Cadmium (ppm) | 109.5792 | 105.7019 | 113.4564 | 107.7087 | 111.4496
Ce, Cerium (ppm) | 36.3202 | 32.0122 | 40.6282 | 35.2806 | 37.3599
Co, Cobalt (ppm) | 25.7742 | 24.1868 | 27.3616 | 25.0352 | 26.5132
Cr, Chromium (ppm) | 20.9877 | 19.392 | 22.5834 | 19.8992 | 22.0763
Cs, Caesium (ppm) | 0.84491 | 0.74611 | 0.94372 | 0.8179 | 0.87193
Cu, Copper (wt.%) | 0.036823 | 0.035736 | 0.03791 | 0.036029 | 0.037617
Fe, Iron (wt.%) | 5.9115 | 5.7605 | 6.0624 | 5.8115 | 6.0115
Ga, Gallium (ppm) | 2.4924 | 2.2778 | 2.7071 | 2.4053 | 2.5795
Ge, Germanium (ppm) | 0.11083 | 0.090091 | 0.13158 | IND | IND
Hf, Hafnium (ppm) | 0.35315 | 0.31678 | 0.38952 | 0.33592 | 0.37037
Hg, Mercury (ppm) | 0.77914 | 0.71734 | 0.84094 | 0.74095 | 0.81733
In, Indium (ppm) | 0.25419 | 0.23432 | 0.27406 | 0.243 | 0.26537
K, Potassium (wt.%) | 0.2455 | 0.233 | 0.258 | 0.23527 | 0.25573
La, Lanthanum (ppm) | 16.8537 | 15.0755 | 18.632 | 16.3905 | 17.317
Li, Lithium (ppm) | 7.0036 | 6.755 | 7.2523 | 6.8199 | 7.1874
Mg, Magnesium (wt.%) | 0.8224 | 0.79226 | 0.85254 | 0.80537 | 0.83943
Mn, Manganese (wt.%) | 0.18452 | 0.17849 | 0.19055 | 0.18119 | 0.18785
Mo, Molybdenum (ppm) | 5.0831 | 4.8064 | 5.3598 | 4.9324 | 5.2338
Nb, Niobium (ppm) | 0.056572 | 0.037204 | 0.075941 | IND | IND
Ni, Nickel (ppm) | 13.8723 | 13.1655 | 14.5792 | 13.2373 | 14.5074
P, Phosphorus (wt.%) | 0.027695 | 0.026778 | 0.028613 | 0.026974 | 0.028417
Pb, Lead (wt.%) | 3.8351 | 3.7597 | 3.9106 | 3.7904 | 3.8799
Rb, Rubidium (ppm) | 14.5261 | 13.3801 | 15.672 | 14.0814 | 14.9707
S, Sulphur (wt.%) | 4.5493 | 4.4207 | 4.6779 | 4.4575 | 4.6411
Sb, Antimony (ppm) | 33.0633 | 30.7633 | 35.3633 | 32.0192 | 34.1074
Sc, Scandium (ppm) | 1.9286 | 1.7915 | 2.0657 | 1.8144 | 2.0428
Se, Selenium (ppm) | 1.9408 | 1.5517 | 2.3299 | 1.7725 | 2.1091
Sn, Tin (ppm) | 0.51544 | 0.43877 | 0.59211 | 0.46827 | 0.56261
Sr, Strontium (ppm) | 18.6217 | 16.9618 | 20.2817 | 17.9747 | 19.2688
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.31448 | 0.27509 | 0.35387 | 0.29227 | 0.33668
Te, Tellurium (ppm) | 0.0685 | 0.050168 | 0.086833 | IND | IND
Th, Thorium (ppm) | 6.3637 | 5.8669 | 6.8605 | 6.1765 | 6.5509
Ti, Titanium (wt.%) | 0.0055248 | 0.0046178 | 0.0064318 | 0.0052208 | 0.0058288
Tl, Thallium (ppm) | 10.9262 | 10.3122 | 11.5401 | 10.6634 | 11.1889
U, Uranium (ppm) | 1.8139 | 1.7037 | 1.9241 | 1.7532 | 1.8745
V, Vanadium (ppm) | 17.7802 | 16.6577 | 18.9027 | 16.8944 | 18.666
W, Tungsten (ppm) | 0.14203 | 0.12027 | 0.16379 | 0.12953 | 0.15453
Y, Yttrium (ppm) | 5.6309 | 5.3352 | 5.9266 | 5.4669 | 5.7949
Yb, Ytterbium (ppm) | 0.55648 | 0.45786 | 0.6551 | 0.52814 | 0.58482
Zn, Zinc (wt.%) | 3.4491 | 3.3914 | 3.5068 | 3.4091 | 3.4891
Zr, Zirconium (ppm) | 11.8453 | 11.2214 | 12.4692 | 11.4536 | 12.237
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 9.5346 | 9.4147 | 9.6546 | 9.4301 | 9.6392
BaO, Barium oxide (wt.%) | 0.26931 | 0.25966 | 0.27896 | 0.26049 | 0.27813
Bi, Bismuth (ppm) | < 100 | IND | IND | IND | IND
CaO, Calcium oxide (wt.%) | 1.913 | 1.8849 | 1.9412 | 1.8883 | 1.9377
Cr2O3, Chromium(III) oxide (ppm) | 150.3948 | 123.2605 | 177.5291 | IND | IND
Cu, Copper (wt.%) | 0.034665 | 0.032166 | 0.037164 | 0.033296 | 0.036034
Fe2O3, Iron(III) oxide (wt.%) | 8.9875 | 8.8791 | 9.0959 | 8.8979 | 9.0771
HfO2, Hafnium dioxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 2.4291 | 2.3987 | 2.4596 | 2.3982 | 2.4601
MgO, Magnesium oxide (wt.%) | 1.7631 | 1.7056 | 1.8206 | 1.7295 | 1.7966
MnO, Manganese oxide (wt.%) | 0.24639 | 0.23851 | 0.25427 | 0.23907 | 0.25371
P2O5, Phosphorus(V) oxide (wt.%) | 0.070906 | 0.062693 | 0.079119 | 0.066155 | 0.075657
Pb, Lead (wt.%) | 3.8597 | 3.7847 | 3.9347 | 3.7999 | 3.9194
S, Sulphur (wt.%) | 4.5904 | 4.4455 | 4.7353 | 4.5183 | 4.6625
SiO2, Silicon dioxide (wt.%) | 57.6097 | 56.941 | 58.2785 | 57.1403 | 58.0791
Sn, Tin (ppm) | < 50 | IND | IND | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.40653 | 0.39307 | 0.41999 | 0.39671 | 0.41635
Zn, Zinc (wt.%) | 3.4686 | 3.3991 | 3.538 | 3.4287 | 3.5085
Thermogravimetry
LOI1000, Loss On Ignition @1000 °C (wt.%) | 7.9435 | 7.7621 | 8.1248 | 7.8007 | 8.0863
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 4.9554 | 4.8846 | 5.0261 | 4.9038 | 5.0069
As, Arsenic (ppm) | 92.6456 | 87.9394 | 97.3517 | 89.5731 | 95.7181
Ba, Barium (wt.%) | 0.23783 | 0.2317 | 0.24396 | 0.23463 | 0.24103
Be, Beryllium (ppm) | 2.3463 | 1.8225 | 2.8701 | 2.1739 | 2.5186
Bi, Bismuth (ppm) | 0.44995 | 0.33887 | 0.56102 | IND | IND
Ca, Calcium (wt.%) | 1.3448 | 1.2845 | 1.405 | 1.3166 | 1.373
Cd, Cadmium (ppm) | 111.6397 | 107.6631 | 115.6163 | 108.6505 | 114.6289
Ce, Cerium (ppm) | 60.4291 | 58.2302 | 62.6281 | 58.3046 | 62.5536
Co, Cobalt (ppm) | 26.4422 | 25.0235 | 27.861 | 25.5597 | 27.3248
Cr, Chromium (ppm) | 78.7981 | 73.8269 | 83.7694 | 76.4083 | 81.188
Cs, Caesium (ppm) | 3.2419 | 3.0831 | 3.4007 | 3.0504 | 3.4333
Cu, Copper (wt.%) | 0.036316 | 0.034938 | 0.037695 | 0.035475 | 0.037158
Dy, Dysprosium (ppm) | 3.573 | 3.2085 | 3.9374 | 3.3518 | 3.7941
Er, Erbium (ppm) | 2.125 | 1.9721 | 2.2779 | 1.9993 | 2.2507
Eu, Europium (ppm) | 1.0036 | 0.91908 | 1.0881 | 0.9333 | 1.0739
Fe, Iron (wt.%) | 6.2266 | 6.1008 | 6.3524 | 6.157 | 6.2961
Ga, Gallium (ppm) | 14.5973 | 13.5199 | 15.6748 | 14.1962 | 14.9985
Gd, Gadolinium (ppm) | 4.3683 | 4.0383 | 4.6983 | 4.1291 | 4.6075
Ge, Germanium (ppm) | 10.4054 | 9.6258 | 11.185 | IND | IND
Hf, Hafnium (ppm) | 3.022 | 2.7306 | 3.3135 | 2.8249 | 3.2192
Ho, Holmium (ppm) | 0.71869 | 0.67382 | 0.76355 | 0.67019 | 0.76719
In, Indium (ppm) | 0.30327 | 0.25723 | 0.34932 | IND | IND
K, Potassium (wt.%) | 2.0466 | 2.0137 | 2.0795 | 2.0157 | 2.0775
La, Lanthanum (ppm) | 31.7944 | 30.6107 | 32.9781 | 30.6005 | 32.9883
Li, Lithium (ppm) | 23.5877 | 22.0734 | 25.1021 | 21.9652 | 25.2103
Lu, Lutetium (ppm) | 0.30467 | 0.29061 | 0.31872 | IND | IND
Mg, Magnesium (wt.%) | 1.0465 | 1.0206 | 1.0725 | 1.0308 | 1.0623
Mn, Manganese (wt.%) | 0.18803 | 0.18236 | 0.1937 | 0.18544 | 0.19061
Mo, Molybdenum (ppm) | 5.1378 | 4.6151 | 5.6605 | IND | IND
Nb, Niobium (ppm) | 7.9674 | 7.0773 | 8.8575 | 7.5626 | 8.3722
Nd, Neodymium (ppm) | 27.4076 | 26.0273 | 28.7879 | 26.1735 | 28.6417
Ni, Nickel (ppm) | 19.305 | 14.5368 | 24.0732 | IND | IND
P, Phosphorus (wt.%) | 0.031288 | 0.028674 | 0.033902 | 0.028784 | 0.033793
Pb, Lead (wt.%) | 3.7712 | 3.693 | 3.8494 | 3.7201 | 3.8223
Pr, Praseodymium (ppm) | 7.4635 | 7.1238 | 7.8033 | 7.1311 | 7.7959
Rb, Rubidium (ppm) | 109.3781 | 105.0735 | 113.6827 | 105.3004 | 113.4557
S, Sulphur (wt.%) | 4.5735 | 4.4437 | 4.7033 | 4.5082 | 4.6388
Sb, Antimony (ppm) | 48.5471 | 44.3104 | 52.7839 | 45.8287 | 51.2655
Sc, Scandium (ppm) | 8.3806 | 7.674 | 9.0871 | IND | IND
Si, Silicon (wt.%) | 27.2539 | 26.5829 | 27.925 | 26.8508 | 27.657
Sm, Samarium (ppm) | 5.1112 | 4.7887 | 5.4337 | 4.8605 | 5.3618
Sn, Tin (ppm) | 2.7796 | 2.0936 | 3.4656 | IND | IND
Sr, Strontium (ppm) | 46.6095 | 43.7026 | 49.5163 | 44.7425 | 48.4764
Ta, Tantalum (ppm) | 0.71618 | 0.56672 | 0.86564 | IND | IND
Tb, Terbium (ppm) | 0.61929 | 0.54073 | 0.69784 | 0.57646 | 0.66212
Th, Thorium (ppm) | 10.6504 | 10.3102 | 10.9905 | 10.3602 | 10.9406
Ti, Titanium (wt.%) | 0.2374 | 0.22964 | 0.24517 | 0.23215 | 0.24265
Tl, Thallium (ppm) | 20.0867 | 18.9504 | 21.223 | 19.4532 | 20.7203
Tm, Thulium (ppm) | 0.30513 | 0.28241 | 0.32784 | 0.27033 | 0.33992
U, Uranium (ppm) | 3.6061 | 3.3668 | 3.8455 | 3.4559 | 3.7563
V, Vanadium (ppm) | 137.148 | 131.7876 | 142.5084 | 133.731 | 140.5649
Y, Yttrium (ppm) | 20.0938 | 19.2366 | 20.9509 | 19.5415 | 20.646
Yb, Ytterbium (ppm) | 2.0567 | 1.8536 | 2.2597 | 1.8794 | 2.2339
Zn, Zinc (wt.%) | 3.4463 | 3.3825 | 3.5102 | 3.4014 | 3.4913
Zr, Zirconium (ppm) | 103.3971 | 95.571 | 111.2232 | 96.8989 | 109.8953
Infrared Combustion
C, Carbon (wt.%) | 1.5855 | 1.5585 | 1.6125 | 1.5704 | 1.6007
S, Sulphur (wt.%) | 4.6628 | 4.5937 | 4.7318 | 4.6145 | 4.7111
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 313
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | < 10 | Hg | ppm | < 1
Ba | ppm | 353.1194 | Se | ppm | 2.6471
Aqua Regia Digestion
Ba | ppm | 44.6718 | Ho | ppm | 0.27288 | Pr | ppm | 4.4319
Dy | ppm | 1.5228 | Lu | ppm | 0.091785 | Pt | ppb | < 5
Er | ppm | 0.66464 | Na | wt.% | 0.011412 | Re | ppm | 0.0025972
Eu | ppm | 0.62352 | Nd | ppm | 17.6571 | Sm | ppm | 3.1512
Gd | ppm | 2.5989 | Pd | ppb | 41 | Tm | ppm | 0.08867
Borate Fusion XRF
Ag | ppm | < 10 | Ho | ppm | < 10 | SrO | ppm | 75.9613
As | ppm | 105 | In | ppm | < 100 | Ta | ppm | 12
Au | ppm | < 10 | La | ppm | 61.7385 | Tb | ppm | < 10
Cd | ppm | 126.3333 | Lu | ppm | < 10 | Te | ppm | < 100
Ce | ppm | 78.483 | Mo | ppm | 40.0833 | Th | ppm | 515.5
Cl | wt.% | 0.30824 | Na2O | wt.% | 0.22187 | Tl | ppm | < 10
Co | ppm | 73.8347 | Nb | ppm | 60.3333 | Tm | ppm | < 10
Cs | ppm | 22.1667 | Nd | ppm | 103 | U | ppm | 12.8333
Dy | ppm | 74.8333 | Ni | ppm | 50.4966 | V | ppm | 127.2209
Er | ppm | < 10 | Pr | ppm | < 10 | W | ppm | 67.7222
Eu | ppm | < 10 | Rb | ppm | 83.3333 | Y | ppm | 54.9879
Ga | ppm | 98.8333 | Sb | ppm | < 50 | Yb | ppm | < 10
Gd | ppm | < 10 | Sc | ppm | 16.8333 | Zr | ppm | 106.25
Ge | ppm | < 10 | Se | ppm | < 10
Hg | ppm | < 100 | Sm | ppm | 24
Borate / Peroxide Fusion ICP
Ag | ppm | 76.7815 | Na | wt.% | 0.093348 | Te | ppm | < 1
B | ppm | 96.1422 | Re | ppm | < 0.1 | W | ppm | 2.022
Hg | ppm | < 5 | Se | ppm | < 10
Laser Ablation ICP-MS
Ag | ppm | 70.3 | Hf | ppm | 2.77 | Sm | ppm | 5.035
As | ppm | 93.6 | Ho | ppm | 0.715 | Sn | ppm | 2.6
Ba | wt.% | 0.236 | In | ppm | 0.25 | Sr | ppm | 44.15
Be | ppm | 2.4 | La | ppm | 31.05 | Ta | ppm | 0.325
Bi | ppm | 0.4 | Lu | ppm | 0.275 | Tb | ppm | 0.605
Ce | ppm | 57.1 | Mn | wt.% | 0.1825 | Te | ppm | 0.15
Co | ppm | 25.1 | Mo | ppm | 5.4 | Th | ppm | 10.25
Cr | ppm | 71 | Nb | ppm | 8.285 | Ti | wt.% | 0.2415
Cs | ppm | 3.085 | Nd | ppm | 27.2 | Tm | ppm | 0.305
Cu | wt.% | 0.036 | Ni | ppm | 15 | U | ppm | 3.475
Dy | ppm | 3.62 | Pb | wt.% | 3.865 | V | ppm | 136.5
Er | ppm | 2.15 | Pr | ppm | 7.36 | W | ppm | 0.5
Eu | ppm | 0.95 | Rb | ppm | 105.5 | Y | ppm | 20.25
Ga | ppm | 13.8 | Re | ppm | < 0.01 | Yb | ppm | 2.105
Gd | ppm | 3.92 | Sb | ppm | 49.5 | Zr | ppm | 101.5
Ge | ppm | 8.525 | Sc | ppm | 8.45
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 313
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 75.573 | 2.1013 | 71.3704 | 79.7755 | 69.2692 | 81.8768 | 2.78% | 5.561% | 8.341% | 71.7943 | 79.3516
Al, wt.% | 4.8992 | 0.20707 | 4.485 | 5.3133 | 4.278 | 5.5204 | 4.227% | 8.453% | 12.68% | 4.6542 | 5.1441
As, ppm | 93.8022 | 3.5895 | 86.6233 | 100.9811 | 83.0338 | 104.5706 | 3.827% | 7.653% | 11.48% | 89.1121 | 98.4923
Be, ppm | 2.1866 | 0.17415 | 1.8383 | 2.5349 | 1.6641 | 2.709 | 7.965% | 15.93% | 23.89% | 2.0772 | 2.2959
Bi, ppm | 0.43355 | 0.02285 | 0.38785 | 0.47925 | 0.365 | 0.5021 | 5.27% | 10.54% | 15.81% | 0.41187 | 0.45522
Ca, wt.% | 1.3547 | 0.03945 | 1.2758 | 1.4336 | 1.2364 | 1.4731 | 2.912% | 5.824% | 8.736% | 1.287 | 1.4224
Cd, ppm | 113.1288 | 5.5437 | 102.0413 | 124.2162 | 96.4976 | 129.7599 | 4.9% | 9.801% | 14.7% | 107.4723 | 118.7852
Ce, ppm | 56.1653 | 4.8636 | 46.4381 | 65.8926 | 41.5744 | 70.7563 | 8.66% | 17.32% | 25.98% | 53.3571 | 58.9736
Co, ppm | 25.5278 | 1.0286 | 23.4705 | 27.5851 | 22.4418 | 28.6137 | 4.03% | 8.059% | 12.09% | 24.2514 | 26.8042
Cr, ppm | 68.9483 | 9.2091 | 50.53 | 87.3666 | 41.3209 | 96.5757 | 13.36% | 26.71% | 40.07% | 65.5009 | 72.3957
Cs, ppm | 3.1218 | 0.12209 | 2.8777 | 3.366 | 2.7556 | 3.4881 | 3.911% | 7.821% | 11.73% | 2.9658 | 3.2779
Cu, wt.% | 0.035683 | 0.0008542 | 0.033974 | 0.037391 | 0.03312 | 0.038245 | 2.394% | 4.788% | 7.182% | 0.033898 | 0.037467
Dy, ppm | 2.6886 | 0.43576 | 1.8171 | 3.5601 | 1.3813 | 3.9959 | 16.21% | 32.42% | 48.62% | 2.5542 | 2.823
Er, ppm | 1.3424 | 0.17397 | 0.99448 | 1.6904 | 0.82051 | 1.8643 | 12.96% | 25.92% | 38.88% | 1.2753 | 1.4095
Eu, ppm | 0.90384 | 0.079251 | 0.74534 | 1.0623 | 0.66609 | 1.1416 | 8.768% | 17.54% | 26.3% | 0.85865 | 0.94903
Fe, wt.% | 6.1545 | 0.19264 | 5.7692 | 6.5398 | 5.5766 | 6.7324 | 3.13% | 6.26% | 9.39% | 5.8468 | 6.4622
Ga, ppm | 14.2576 | 1.0178 | 12.222 | 16.2933 | 11.2041 | 17.3111 | 7.139% | 14.28% | 21.42% | 13.5447 | 14.9705
Gd, ppm | 3.8569 | 0.57071 | 2.7155 | 4.9983 | 2.1447 | 5.569 | 14.8% | 29.59% | 44.39% | 3.664 | 4.0497
Ge, ppm | 0.13559 | 0.026201 | 0.083189 | 0.18799 | 0.056988 | 0.21419 | 19.32% | 38.65% | 57.97% | 0.12881 | 0.14237
Hf, ppm | 2.0754 | 0.15045 | 1.7745 | 2.3763 | 1.6241 | 2.5268 | 7.249% | 14.5% | 21.75% | 1.9716 | 2.1792
Ho, ppm | 0.49743 | 0.02461 | 0.44821 | 0.54665 | 0.4236 | 0.57126 | 4.947% | 9.895% | 14.84% | 0.47256 | 0.5223
In, ppm | 0.29068 | 0.019872 | 0.25094 | 0.33042 | 0.23106 | 0.35029 | 6.836% | 13.67% | 20.51% | 0.27615 | 0.30521
K, wt.% | 1.9987 | 0.064773 | 1.8691 | 2.1282 | 1.8044 | 2.193 | 3.241% | 6.482% | 9.722% | 1.8987 | 2.0986
La, ppm | 28.8871 | 3.2513 | 22.3845 | 35.3896 | 19.1332 | 38.6409 | 11.26% | 22.51% | 33.77% | 27.4427 | 30.3314
Li, ppm | 23.4667 | 1.2068 | 21.0531 | 25.8803 | 19.8463 | 27.0871 | 5.143% | 10.29% | 15.43% | 22.2933 | 24.64
Lu, ppm | 0.21098 | 0.042359 | 0.12626 | 0.2957 | 0.083903 | 0.33806 | 20.08% | 40.15% | 60.23% | 0.20043 | 0.22153
Mg, wt.% | 1.0263 | 0.039973 | 0.94633 | 1.1062 | 0.90635 | 1.1462 | 3.895% | 7.79% | 11.68% | 0.97496 | 1.0776
Mn, wt.% | 0.18567 | 0.0073449 | 0.17098 | 0.20036 | 0.16363 | 0.2077 | 3.956% | 7.912% | 11.87% | 0.17638 | 0.19495
Mo, ppm | 5.2571 | 0.29983 | 4.6574 | 5.8568 | 4.3576 | 6.1566 | 5.703% | 11.41% | 17.11% | 4.9942 | 5.52
Na, wt.% | 0.075703 | 0.0061997 | 0.063304 | 0.088103 | 0.057104 | 0.094302 | 8.189% | 16.38% | 24.57% | 0.071918 | 0.079488
Nb, ppm | 3.3011 | 0.44345 | 2.4142 | 4.1881 | 1.9708 | 4.6315 | 13.43% | 26.87% | 40.3% | 3.1361 | 3.4662
Nd, ppm | 26.5575 | 1.5568 | 23.4439 | 29.6712 | 21.887 | 31.228 | 5.862% | 11.72% | 17.59% | 25.2296 | 27.8854
Ni, ppm | 15.9264 | 0.98295 | 13.9605 | 17.8923 | 12.9775 | 18.8752 | 6.172% | 12.34% | 18.52% | 15.13 | 16.7227
P, wt.% | 0.031266 | 0.0012661 | 0.028734 | 0.033799 | 0.027468 | 0.035065 | 4.049% | 8.099% | 12.15% | 0.029703 | 0.03283
Pb, wt.% | 3.8963 | 0.10978 | 3.6767 | 4.1158 | 3.5669 | 4.2256 | 2.818% | 5.635% | 8.453% | 3.7015 | 4.0911
Pr, ppm | 6.831 | 0.73846 | 5.3541 | 8.308 | 4.6157 | 9.0464 | 10.81% | 21.62% | 32.43% | 6.4895 | 7.1726
Rb, ppm | 107.6525 | 5.3301 | 96.9924 | 118.3127 | 91.6623 | 123.6427 | 4.951% | 9.902% | 14.85% | 102.2699 | 113.0351
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 4.4423 | 0.21849 | 4.0053 | 4.8792 | 3.7868 | 5.0977 | 4.919% | 9.837% | 14.76% | 4.2201 | 4.6644
Sb, ppm | 44.4058 | 2.8041 | 38.7975 | 50.014 | 35.9934 | 52.8181 | 6.315% | 12.63% | 18.94% | 42.1855 | 46.6261
Sc, ppm | 8.7012 | 0.59099 | 7.5192 | 9.8832 | 6.9282 | 10.4742 | 6.792% | 13.58% | 20.38% | 8.2662 | 9.1363
Sm, ppm | 4.8749 | 0.2983 | 4.2783 | 5.4715 | 3.98 | 5.7698 | 6.119% | 12.24% | 18.36% | 4.6312 | 5.1187
Sn, ppm | 2.5852 | 0.21638 | 2.1525 | 3.018 | 1.9361 | 3.2344 | 8.37% | 16.74% | 25.11% | 2.456 | 2.7145
Sr, ppm | 41.676 | 1.8992 | 37.8777 | 45.4744 | 35.9785 | 47.3736 | 4.557% | 9.114% | 13.67% | 39.5922 | 43.7598
Ta, ppm | 0.2391 | 0.066129 | 0.10685 | 0.37136 | 0.040717 | 0.43749 | 27.66% | 55.31% | 82.97% | 0.22715 | 0.25106
Tb, ppm | 0.50888 | 0.066296 | 0.37629 | 0.64148 | 0.31 | 0.70777 | 13.03% | 26.06% | 39.08% | 0.48344 | 0.53433
Te, ppm | 0.084432 | 0.020154 | 0.044124 | 0.12474 | 0.02397 | 0.14489 | 23.87% | 47.74% | 71.61% | 0.08021 | 0.088653
Th, ppm | 9.9832 | 0.78217 | 8.4188 | 11.5475 | 7.6366 | 12.3297 | 7.835% | 15.67% | 23.5% | 9.484 | 10.4823
Ti, wt.% | 0.15459 | 0.010839 | 0.13291 | 0.17626 | 0.12207 | 0.1871 | 7.012% | 14.02% | 21.03% | 0.14686 | 0.16232
Tl, ppm | 20.0787 | 1.2007 | 17.6774 | 22.48 | 16.4767 | 23.6807 | 5.98% | 11.96% | 17.94% | 19.0747 | 21.0826
Tm, ppm | 0.17925 | 0.033511 | 0.11222 | 0.24627 | 0.078714 | 0.27978 | 18.7% | 37.39% | 56.09% | 0.17029 | 0.18821
U, ppm | 3.4476 | 0.25558 | 2.9364 | 3.9587 | 2.6808 | 4.2143 | 7.413% | 14.83% | 22.24% | 3.2752 | 3.6199
V, ppm | 133.2304 | 4.4143 | 124.4018 | 142.0589 | 119.9875 | 146.4732 | 3.313% | 6.627% | 9.94% | 126.5689 | 139.8919
W, ppm | 1.3698 | 0.20073 | 0.96834 | 1.7712 | 0.76761 | 1.972 | 14.65% | 29.31% | 43.96% | 1.3013 | 1.4383
Y, ppm | 13.499 | 1.5982 | 10.3025 | 16.6954 | 8.7043 | 18.2937 | 11.84% | 23.68% | 35.52% | 12.824 | 14.1739
Yb, ppm | 1.4549 | 0.28077 | 0.89332 | 2.0164 | 0.61255 | 2.2971 | 19.3% | 38.6% | 57.9% | 1.3821 | 1.5276
Zn, wt.% | 3.5482 | 0.090454 | 3.3673 | 3.7291 | 3.2768 | 3.8195 | 2.549% | 5.099% | 7.648% | 3.3708 | 3.7256
Zr, ppm | 69.8741 | 4.1005 | 61.6731 | 78.0751 | 57.5727 | 82.1756 | 5.868% | 11.74% | 17.61% | 66.3804 | 73.3678
Aqua Regia Digestion
Ag, ppm | 74.5965 | 2.5197 | 69.5571 | 79.6359 | 67.0374 | 82.1556 | 3.378% | 6.756% | 10.13% | 70.8667 | 78.3263
Al, wt.% | 0.78732 | 0.056957 | 0.6734 | 0.90123 | 0.61644 | 0.95819 | 7.234% | 14.47% | 21.7% | 0.74795 | 0.82668
As, ppm | 93.3309 | 7.5321 | 78.2668 | 108.395 | 70.7348 | 115.9271 | 8.07% | 16.14% | 24.21% | 88.6644 | 97.9975
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.5098 | 0.038885 | 0.43203 | 0.58757 | 0.39314 | 0.62645 | 7.627% | 15.25% | 22.88% | 0.48431 | 0.53529
Bi, ppm | 0.31715 | 0.028098 | 0.26096 | 0.37335 | 0.23286 | 0.40145 | 8.86% | 17.72% | 26.58% | 0.3013 | 0.33301
Ca, wt.% | 1.3414 | 0.052624 | 1.2361 | 1.4466 | 1.1835 | 1.4993 | 3.923% | 7.846% | 11.77% | 1.2743 | 1.4084
Cd, ppm | 109.5792 | 6.3418 | 96.8956 | 122.2628 | 90.5538 | 128.6046 | 5.787% | 11.57% | 17.36% | 104.1002 | 115.0581
Ce, ppm | 36.3202 | 8.3401 | 19.6401 | 53.0004 | 11.3001 | 61.3404 | 22.96% | 45.93% | 68.89% | 34.5042 | 38.1363
Co, ppm | 25.7742 | 2.1311 | 21.5121 | 30.0363 | 19.381 | 32.1674 | 8.268% | 16.54% | 24.8% | 24.4855 | 27.0629
Cr, ppm | 20.9877 | 1.9557 | 17.0763 | 24.8992 | 15.1205 | 26.8549 | 9.318% | 18.64% | 27.96% | 19.9383 | 22.0371
Cs, ppm | 0.84491 | 0.14669 | 0.55153 | 1.1383 | 0.40484 | 1.285 | 17.36% | 34.72% | 52.09% | 0.80267 | 0.88716
Cu, wt.% | 0.036823 | 0.0014374 | 0.033948 | 0.039698 | 0.032511 | 0.041135 | 3.903% | 7.807% | 11.71% | 0.034982 | 0.038664
Fe, wt.% | 5.9115 | 0.22603 | 5.4594 | 6.3635 | 5.2334 | 6.5896 | 3.824% | 7.647% | 11.47% | 5.6159 | 6.207
Ga, ppm | 2.4924 | 0.29334 | 1.9058 | 3.0791 | 1.6124 | 3.3725 | 11.77% | 23.54% | 35.31% | 2.3678 | 2.6171
Ge, ppm | 0.11083 | 0.014614 | 0.081605 | 0.14006 | 0.066991 | 0.15468 | 13.19% | 26.37% | 39.56% | 0.10529 | 0.11638
Hf, ppm | 0.35315 | 0.038168 | 0.27681 | 0.42948 | 0.23864 | 0.46765 | 10.81% | 21.62% | 32.42% | 0.33549 | 0.3708
Hg, ppm | 0.77914 | 0.051954 | 0.67523 | 0.88305 | 0.62328 | 0.935 | 6.668% | 13.34% | 20% | 0.74018 | 0.8181
In, ppm | 0.25419 | 0.01814 | 0.21791 | 0.29047 | 0.19977 | 0.30861 | 7.136% | 14.27% | 21.41% | 0.24148 | 0.2669
K, wt.% | 0.2455 | 0.020995 | 0.20351 | 0.28749 | 0.18251 | 0.30848 | 8.552% | 17.1% | 25.66% | 0.23322 | 0.25777
La, ppm | 16.8537 | 3.3074 | 10.239 | 23.4685 | 6.9316 | 26.7758 | 19.62% | 39.25% | 58.87% | 16.011 | 17.6964
Li, ppm | 7.0036 | 0.31737 | 6.3689 | 7.6384 | 6.0515 | 7.9557 | 4.531% | 9.063% | 13.59% | 6.6534 | 7.3538
Mg, wt.% | 0.8224 | 0.040306 | 0.74179 | 0.90301 | 0.70148 | 0.94332 | 4.901% | 9.802% | 14.7% | 0.78128 | 0.86352
Mn, wt.% | 0.18452 | 0.00806 | 0.1684 | 0.20064 | 0.16034 | 0.2087 | 4.368% | 8.736% | 13.1% | 0.17529 | 0.19375
Mo, ppm | 5.0831 | 0.28181 | 4.5195 | 5.6467 | 4.2377 | 5.9285 | 5.544% | 11.09% | 16.63% | 4.8289 | 5.3373
Nb, ppm | 0.056572 | 0.010649 | 0.035275 | 0.07787 | 0.024626 | 0.088519 | 18.82% | 37.65% | 56.47% | 0.053744 | 0.059401
Ni, ppm | 13.8723 | 0.90854 | 12.0552 | 15.6894 | 11.1467 | 16.598 | 6.549% | 13.1% | 19.65% | 13.1787 | 14.5659
P, wt.% | 0.027695 | 0.0014157 | 0.024864 | 0.030527 | 0.023448 | 0.031942 | 5.112% | 10.22% | 15.33% | 0.026311 | 0.02908
Pb, wt.% | 3.8351 | 0.10851 | 3.6181 | 4.0522 | 3.5096 | 4.1607 | 2.829% | 5.659% | 8.488% | 3.6434 | 4.0269
Rb, ppm | 14.5261 | 1.8023 | 10.9216 | 18.1306 | 9.1193 | 19.9328 | 12.41% | 24.81% | 37.22% | 13.7998 | 15.2524
S, wt.% | 4.5493 | 0.14818 | 4.2529 | 4.8456 | 4.1047 | 4.9938 | 3.257% | 6.514% | 9.772% | 4.3218 | 4.7767
Sb, ppm | 33.0633 | 3.2345 | 26.5943 | 39.5323 | 23.3598 | 42.7668 | 9.783% | 19.57% | 29.35% | 31.4101 | 34.7165
Sc, ppm | 1.9286 | 0.13206 | 1.6645 | 2.1927 | 1.5324 | 2.3248 | 6.847% | 13.69% | 20.54% | 1.8322 | 2.025
Se, ppm | 1.9408 | 0.44579 | 1.0493 | 2.8324 | 0.60347 | 3.2782 | 22.97% | 45.94% | 68.91% | 1.8438 | 2.0379
Sn, ppm | 0.51544 | 0.073244 | 0.36895 | 0.66193 | 0.29571 | 0.73517 | 14.21% | 28.42% | 42.63% | 0.48967 | 0.54121
Sr, ppm | 18.6217 | 2.7143 | 13.1932 | 24.0503 | 10.4789 | 26.7646 | 14.58% | 29.15% | 43.73% | 17.6907 | 19.5528
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.31448 | 0.031125 | 0.25223 | 0.37673 | 0.2211 | 0.40785 | 9.897% | 19.79% | 29.69% | 0.29875 | 0.3302
Te, ppm | 0.0685 | 0.013402 | 0.041697 | 0.095304 | 0.028295 | 0.10871 | 19.56% | 39.13% | 58.69% | 0.065075 | 0.071925
Th, ppm | 6.3637 | 0.92607 | 4.5116 | 8.2159 | 3.5855 | 9.1419 | 14.55% | 29.1% | 43.66% | 6.0455 | 6.6819
Ti, wt.% | 0.0055248 | 0.00080201 | 0.0039208 | 0.0071288 | 0.0031188 | 0.0079308 | 14.52% | 29.03% | 43.55% | 0.0052486 | 0.005801
Tl, ppm | 10.9262 | 0.7499 | 9.4264 | 12.426 | 8.6765 | 13.1759 | 6.863% | 13.73% | 20.59% | 10.3799 | 11.4725
U, ppm | 1.8139 | 0.16029 | 1.4933 | 2.1345 | 1.333 | 2.2948 | 8.837% | 17.67% | 26.51% | 1.7232 | 1.9046
V, ppm | 17.7802 | 1.7027 | 14.3748 | 21.1856 | 12.6721 | 22.8883 | 9.576% | 19.15% | 28.73% | 16.8912 | 18.6692
W, ppm | 0.14203 | 0.018471 | 0.10509 | 0.17897 | 0.086618 | 0.19744 | 13% | 26.01% | 39.01% | 0.13493 | 0.14913
Y, ppm | 5.6309 | 0.3982 | 4.8345 | 6.4273 | 4.4363 | 6.8255 | 7.072% | 14.14% | 21.22% | 5.3494 | 5.9124
Yb, ppm | 0.55648 | 0.043074 | 0.47033 | 0.64263 | 0.42726 | 0.6857 | 7.74% | 15.48% | 23.22% | 0.52866 | 0.5843
Zn, wt.% | 3.4491 | 0.059667 | 3.3298 | 3.5684 | 3.2701 | 3.6281 | 1.73% | 3.46% | 5.19% | 3.2767 | 3.6216
Zr, ppm | 11.8453 | 1.1319 | 9.5815 | 14.1091 | 8.4495 | 15.2411 | 9.556% | 19.11% | 28.67% | 11.253 | 12.4376
Borate Fusion XRF
Al2O3, wt.% | 9.5346 | 0.11393 | 9.3068 | 9.7625 | 9.1929 | 9.8764 | 1.195% | 2.39% | 3.585% | 9.0579 | 10.0114
BaO, wt.% | 0.26931 | 0.010687 | 0.24794 | 0.29068 | 0.23725 | 0.30137 | 3.968% | 7.937% | 11.9% | 0.25584 | 0.28278
Bi, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
CaO, wt.% | 1.913 | 0.036156 | 1.8407 | 1.9853 | 1.8046 | 2.0215 | 1.89% | 3.78% | 5.67% | 1.8174 | 2.0087
Cr2O3, ppm | 150.3948 | 20.6742 | 109.0464 | 191.7433 | 88.3721 | 212.4175 | 13.75% | 27.49% | 41.24% | 142.8751 | 157.9145
Cu, wt.% | 0.034665 | 0.0036192 | 0.027427 | 0.041903 | 0.023807 | 0.045523 | 10.44% | 20.88% | 31.32% | 0.032932 | 0.036398
Fe2O3, wt.% | 8.9875 | 0.1061 | 8.7753 | 9.1997 | 8.6692 | 9.3058 | 1.181% | 2.361% | 3.542% | 8.5381 | 9.4369
HfO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 2.4291 | 0.028949 | 2.3712 | 2.487 | 2.3423 | 2.516 | 1.192% | 2.384% | 3.575% | 2.3077 | 2.5506
MgO, wt.% | 1.7631 | 0.087112 | 1.5888 | 1.9373 | 1.5017 | 2.0244 | 4.941% | 9.882% | 14.82% | 1.6749 | 1.8512
MnO, wt.% | 0.24639 | 0.0062442 | 0.2339 | 0.25888 | 0.22766 | 0.26512 | 2.534% | 5.069% | 7.603% | 0.23407 | 0.25871
P2O5, wt.% | 0.070906 | 0.0062666 | 0.058373 | 0.083439 | 0.052106 | 0.089706 | 8.838% | 17.68% | 26.51% | 0.067361 | 0.074451
Pb, wt.% | 3.8597 | 0.0875 | 3.6847 | 4.0347 | 3.5972 | 4.1222 | 2.267% | 4.534% | 6.801% | 3.6667 | 4.0526
S, wt.% | 4.5904 | 0.13673 | 4.3169 | 4.8638 | 4.1802 | 5.0006 | 2.979% | 5.957% | 8.936% | 4.3609 | 4.8199
SiO2, wt.% | 57.6097 | 0.55116 | 56.5074 | 58.712 | 55.9562 | 59.2632 | 0.9567% | 1.913% | 2.87% | 54.7292 | 60.4902
Sn, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
TiO2, wt.% | 0.40653 | 0.014166 | 0.3782 | 0.43486 | 0.36403 | 0.44903 | 3.485% | 6.969% | 10.45% | 0.3862 | 0.42686
Zn, wt.% | 3.4686 | 0.086957 | 3.2947 | 3.6425 | 3.2077 | 3.7295 | 2.507% | 5.014% | 7.521% | 3.2951 | 3.642
Thermogravimetry
LOI1000, wt.% | 7.9435 | 0.26691 | 7.4097 | 8.4773 | 7.1428 | 8.7442 | 3.36% | 6.72% | 10.08% | 7.5463 | 8.3406
Borate / Peroxide Fusion ICP
Al, wt.% | 4.9554 | 0.07996 | 4.7954 | 5.1153 | 4.7155 | 5.1952 | 1.614% | 3.227% | 4.841% | 4.7076 | 5.2031
As, ppm | 92.6456 | 4.5386 | 83.5685 | 101.7227 | 79.0299 | 106.2612 | 4.899% | 9.798% | 14.7% | 88.0133 | 97.2779
Ba, wt.% | 0.23783 | 0.0063393 | 0.22515 | 0.25051 | 0.21881 | 0.25685 | 2.665% | 5.331% | 7.996% | 0.22594 | 0.24972
Be, ppm | 2.3463 | 0.56995 | 1.2064 | 3.4862 | 0.63641 | 4.0561 | 24.29% | 48.58% | 72.88% | 2.229 | 2.4636
Bi, ppm | 0.44995 | 0.055975 | 0.338 | 0.5619 | 0.28202 | 0.61787 | 12.44% | 24.88% | 37.32% | 0.42745 | 0.47244
Ca, wt.% | 1.3448 | 0.067905 | 1.209 | 1.4806 | 1.1411 | 1.5485 | 5.049% | 10.1% | 15.15% | 1.2775 | 1.412
Cd, ppm | 111.6397 | 2.682 | 106.2758 | 117.0036 | 103.5938 | 119.6856 | 2.402% | 4.805% | 7.207% | 106.0577 | 117.2217
Ce, ppm | 60.4291 | 2.1644 | 56.1002 | 64.758 | 53.9358 | 66.9224 | 3.582% | 7.164% | 10.75% | 57.4077 | 63.4506
Co, ppm | 26.4422 | 1.6881 | 23.0661 | 29.8184 | 21.378 | 31.5064 | 6.384% | 12.77% | 19.15% | 25.1201 | 27.7643
Cr, ppm | 78.7981 | 6.4391 | 65.9199 | 91.6763 | 59.4808 | 98.1154 | 8.172% | 16.34% | 24.51% | 74.8582 | 82.738
Cs, ppm | 3.2419 | 0.15725 | 2.9274 | 3.5564 | 2.7701 | 3.7136 | 4.85% | 9.701% | 14.55% | 3.0798 | 3.404
Cu, wt.% | 0.036316 | 0.0017645 | 0.032787 | 0.039845 | 0.031023 | 0.04161 | 4.859% | 9.717% | 14.58% | 0.0345 | 0.038132
Dy, ppm | 3.573 | 0.22144 | 3.1301 | 4.0158 | 2.9086 | 4.2373 | 6.198% | 12.4% | 18.59% | 3.3943 | 3.7516
Er, ppm | 2.125 | 0.1252 | 1.8746 | 2.3754 | 1.7494 | 2.5006 | 5.892% | 11.78% | 17.67% | 2.0187 | 2.2313
Eu, ppm | 1.0036 | 0.088595 | 0.82642 | 1.1808 | 0.73783 | 1.2694 | 8.828% | 17.66% | 26.48% | 0.95343 | 1.0538
Fe, wt.% | 6.2266 | 0.16768 | 5.8912 | 6.562 | 5.7235 | 6.7296 | 2.693% | 5.386% | 8.079% | 5.9153 | 6.5379
Ga, ppm | 14.5973 | 0.82429 | 12.9488 | 16.2459 | 12.1245 | 17.0702 | 5.647% | 11.29% | 16.94% | 13.8675 | 15.3272
Gd, ppm | 4.3683 | 0.2448 | 3.8787 | 4.8579 | 3.6339 | 5.1027 | 5.604% | 11.21% | 16.81% | 4.1499 | 4.5867
Ge, ppm | 10.4054 | 0.5206 | 9.3642 | 11.4466 | 8.8436 | 11.9672 | 5.003% | 10.01% | 15.01% | 9.8851 | 10.9257
Hf, ppm | 3.022 | 0.2014 | 2.6192 | 3.4248 | 2.4178 | 3.6262 | 6.664% | 13.33% | 19.99% | 2.8709 | 3.1731
Ho, ppm | 0.71869 | 0.021891 | 0.67491 | 0.76247 | 0.65302 | 0.78436 | 3.046% | 6.092% | 9.138% | 0.68275 | 0.75462
In, ppm | 0.30327 | 0.01984 | 0.26359 | 0.34295 | 0.24375 | 0.36279 | 6.542% | 13.08% | 19.63% | 0.28811 | 0.31844
K, wt.% | 2.0466 | 0.035795 | 1.975 | 2.1182 | 1.9392 | 2.154 | 1.749% | 3.498% | 5.247% | 1.9443 | 2.1489
La, ppm | 31.7944 | 1.5886 | 28.6172 | 34.9716 | 27.0286 | 36.5602 | 4.996% | 9.993% | 14.99% | 30.2047 | 33.3841
Li, ppm | 23.5877 | 1.6157 | 20.3563 | 26.8192 | 18.7406 | 28.4349 | 6.85% | 13.7% | 20.55% | 22.4083 | 24.7671
Lu, ppm | 0.30467 | 0.015025 | 0.27462 | 0.33472 | 0.25959 | 0.34974 | 4.932% | 9.863% | 14.79% | 0.28943 | 0.3199
Mg, wt.% | 1.0465 | 0.027364 | 0.99181 | 1.1013 | 0.96444 | 1.1286 | 2.615% | 5.229% | 7.844% | 0.99421 | 1.0989
Mn, wt.% | 0.18803 | 0.0070426 | 0.17394 | 0.20211 | 0.1669 | 0.20916 | 3.746% | 7.491% | 11.24% | 0.17863 | 0.19743
Mo, ppm | 5.1378 | 0.5972 | 3.9434 | 6.3322 | 3.3462 | 6.9294 | 11.62% | 23.25% | 34.87% | 4.8809 | 5.3947
Nb, ppm | 7.9674 | 0.62584 | 6.7158 | 9.2191 | 6.0899 | 9.845 | 7.855% | 15.71% | 23.56% | 7.5691 | 8.3658
Nd, ppm | 27.4076 | 1.7334 | 23.9408 | 30.8744 | 22.2073 | 32.6079 | 6.325% | 12.65% | 18.97% | 26.0372 | 28.778
Ni, ppm | 19.305 | 3.3023 | 12.7003 | 25.9097 | 9.398 | 29.212 | 17.11% | 34.21% | 51.32% | 18.3397 | 20.2703
P, wt.% | 0.031288 | 0.0029255 | 0.025437 | 0.037139 | 0.022512 | 0.040065 | 9.35% | 18.7% | 28.05% | 0.029724 | 0.032853
Pb, wt.% | 3.7712 | 0.075815 | 3.6196 | 3.9228 | 3.5437 | 3.9986 | 2.01% | 4.021% | 6.031% | 3.5826 | 3.9597
Pr, ppm | 7.4635 | 0.33994 | 6.7837 | 8.1434 | 6.4437 | 8.4833 | 4.555% | 9.109% | 13.66% | 7.0904 | 7.8367
Rb, ppm | 109.3781 | 4.9761 | 99.4259 | 119.3303 | 94.4498 | 124.3064 | 4.549% | 9.099% | 13.65% | 103.9092 | 114.847
S, wt.% | 4.5735 | 0.083098 | 4.4073 | 4.7397 | 4.3242 | 4.8228 | 1.817% | 3.634% | 5.451% | 4.3448 | 4.8022
Sb, ppm | 48.5471 | 3.3637 | 41.8197 | 55.2746 | 38.456 | 58.6383 | 6.929% | 13.86% | 20.79% | 46.1198 | 50.9745
Sc, ppm | 8.3806 | 0.6684 | 7.0438 | 9.7173 | 6.3754 | 10.3857 | 7.976% | 15.95% | 23.93% | 7.9615 | 8.7996
Si, wt.% | 27.2539 | 0.58662 | 26.0807 | 28.4272 | 25.4941 | 29.0138 | 2.152% | 4.305% | 6.457% | 25.8912 | 28.6166
Sm, ppm | 5.1112 | 0.31783 | 4.4755 | 5.7468 | 4.1577 | 6.0647 | 6.218% | 12.44% | 18.66% | 4.8556 | 5.3667
Sn, ppm | 2.7796 | 0.29241 | 2.1948 | 3.3644 | 1.9024 | 3.6568 | 10.52% | 21.04% | 31.56% | 2.6406 | 2.9186
Sr, ppm | 46.6095 | 3.8633 | 38.8828 | 54.3361 | 35.0195 | 58.1994 | 8.289% | 16.58% | 24.87% | 44.279 | 48.9399
Ta, ppm | 0.71618 | 0.079498 | 0.55718 | 0.87518 | 0.47769 | 0.95467 | 11.1% | 22.2% | 33.3% | 0.68037 | 0.75199
Tb, ppm | 0.61929 | 0.043862 | 0.53156 | 0.70701 | 0.4877 | 0.75087 | 7.083% | 14.17% | 21.25% | 0.58832 | 0.65025
Th, ppm | 10.6504 | 0.46134 | 9.7277 | 11.5731 | 9.2664 | 12.0344 | 4.332% | 8.663% | 13% | 10.1179 | 11.1829
Ti, wt.% | 0.2374 | 0.0077948 | 0.22181 | 0.25299 | 0.21402 | 0.26079 | 3.283% | 6.567% | 9.85% | 0.22553 | 0.24927
Tl, ppm | 20.0867 | 0.85093 | 18.3849 | 21.7886 | 17.5339 | 22.6395 | 4.236% | 8.473% | 12.71% | 19.0824 | 21.0911
Tm, ppm | 0.30513 | 0.0099135 | 0.2853 | 0.32495 | 0.27539 | 0.33487 | 3.249% | 6.498% | 9.747% | 0.28987 | 0.32038
U, ppm | 3.6061 | 0.14468 | 3.3168 | 3.8955 | 3.1721 | 4.0402 | 4.012% | 8.024% | 12.04% | 3.4258 | 3.7864
V, ppm | 137.148 | 5.3106 | 126.5268 | 147.7692 | 121.2162 | 153.0797 | 3.872% | 7.744% | 11.62% | 130.2906 | 144.0054
Y, ppm | 20.0938 | 0.74372 | 18.6063 | 21.5812 | 17.8626 | 22.3249 | 3.701% | 7.403% | 11.1% | 19.0891 | 21.0984
Yb, ppm | 2.0567 | 0.075199 | 1.9063 | 2.2071 | 1.8311 | 2.2823 | 3.656% | 7.313% | 10.97% | 1.9538 | 2.1595
Zn, wt.% | 3.4463 | 0.090454 | 3.2654 | 3.6272 | 3.175 | 3.7177 | 2.625% | 5.249% | 7.874% | 3.274 | 3.6186
Zr, ppm | 103.3971 | 4.3727 | 94.6517 | 112.1425 | 90.2791 | 116.5152 | 4.229% | 8.458% | 12.69% | 98.2273 | 108.567
Infrared Combustion
C, wt.% | 1.5855 | 0.051562 | 1.4824 | 1.6887 | 1.4308 | 1.7402 | 3.252% | 6.504% | 9.756% | 1.5063 | 1.6648
S, wt.% | 4.6628 | 0.10344 | 4.4559 | 4.8697 | 4.3525 | 4.9731 | 2.219% | 4.437% | 6.656% | 4.4296 | 4.8959
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 313
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PSF*XRF | pyro-sulfate fusion with X-ray fluorescence spectroscopy
Text Values:
- | Not Analysed (Lab 2.17)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 313
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Genalysis, Adelaide, SA, Australia
MSALABS, Vancouver, BC, Canada
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 313 (Certified Value 75.6 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 70.9 | 76.8645 | 75.258 | 76.41 | 71.89 | 75.7531 | 78.9 | 73.9 | 75 | 75.9 | 70.8 | 76.2 | 76.14 | 75.2538 | 75.7 | 75.27 | 76.622 | 73.5 | 77.3 | 77 | 78.7 | 73.3 | 78.38 | 75.18 | 65.8 | 1
| 1 | 2 | 72 | 75.0636 | 75.258 | 76.44 | 71.16 | 76.1969 | 79.1 | 74.6 | 74.8 | 77.3 | 71.7 | 76.1 | 76.27 | 76.2629 | 77.1 | 77.9 | 77.275 | 73.5 | 75.7 | 77 | 79.6 | 72.7 | 78.1 | 74.6 | 67.1 | 10
| 1 | 3 | 72.2 | 77.3169 | 74.919 | 76.9 | 72.54 | 74.549 | 78.4 | 77.1 | 75 | 76.8 | 72.6 | 75.5 | 78.84 | 75.1331 | 76.7 | 77.76 | 78.05 | 72.4 | 76.6 | 77 | 77.5 | 74.8 | 79 | 73.83 | 68 | 16
| 1 | 4 | 71.7 | 77.167 | 75.936 | 75.96 | 74.42 | 75.6998 | 78.1 | 73.9 | 74.6 | 78.4 | 71.9 | 75.8 | 76.19 | 75.8041 | 73.5 | 75.43 | 76.092 | 72.3 | 76.5 | 77 | 77.9 | 74 | 78.72 | 74.95 | 68.9 | 75.573
| 1 | 5 | 72.2 | 76.8993 | 76.162 | 75.76 | 72.2 | 74.6655 | 79.7 | 73.1 | 75.7 | 76.6 | 73 | 75.6 | 77.19 | 76.5224 | 74.7 | 78.44 | 77.806 | 74.4 | 71.4 | 77 | 76.2 | 73.2 | 79.27 | 73.92 | 64.7 | 7
| 1 | 6 | 71.9 | 75.9075 | 73.6421 | 76.77 | 72.96 | 76.2478 | 78.8 | 75.3 | 75.6 | 75.3 | 72.1 | 75.1 | 76.82 | 76.261 | 74.7 | 77.13 | 78.622 | 72.6 | 75.5 | 76 | 79.6 | 73.3 | 77.21 | 74.25 | 66.1
| Mean |  | 71.8167 | 76.5365 | 75.1958 | 76.3733 | 72.5283 | 75.5187 | 78.8333 | 74.65 | 75.1167 | 76.7167 | 72.0167 | 75.7167 | 76.9083 | 75.8729 | 75.4 | 76.9883 | 77.4112 | 73.1167 | 75.5 | 76.8333 | 78.25 | 73.55 | 78.4467 | 74.455 | 66.7667
| Median |  | 71.95 | 76.8819 | 75.258 | 76.425 | 72.37 | 75.7264 | 78.85 | 74.25 | 75 | 76.7 | 72 | 75.7 | 76.545 | 76.0326 | 75.2 | 77.445 | 77.5405 | 73.05 | 76.1 | 77 | 78.3 | 73.3 | 78.55 | 74.425 | 66.6
| Std Dev. |  | 0.48751 | 0.87302 | 0.89216 | 0.44442 | 1.1093 | 0.74126 | 0.55737 | 1.411 | 0.44008 | 1.0834 | 0.76267 | 0.40702 | 1.0327 | 0.57618 | 1.3609 | 1.3368 | 0.93932 | 0.82321 | 2.1119 | 0.40825 | 1.3217 | 0.73959 | 0.73658 | 0.54965 | 1.5384
| Rel.Std.Dev. |  | 0.6788% | 1.141% | 1.186% | 0.5819% | 1.529% | 0.9816% | 0.707% | 1.89% | 0.5859% | 1.412% | 1.059% | 0.5376% | 1.343% | 0.7594% | 1.805% | 1.736% | 1.213% | 1.126% | 2.797% | 0.5313% | 1.689% | 1.006% | 0.939% | 0.7382% | 2.304%
| PDM3 |  | -4.97% | 1.275% | -0.499% | 1.059% | -4.029% | -0.07187% | 4.314% | -1.221% | -0.6038% | 1.513% | -4.706% | 0.1901% | 1.767% | 0.3968% | -0.2289% | 1.873% | 2.432% | -3.25% | -0.09658% | 1.668% | 3.542% | -2.677% | 3.803% | -1.479% | -11.65%
| Z-Score (Absolute) |  | 1.9 | 0.52 | 0.17 | 0.44 | 1.53 | 0 | 1.7 | 0.45 | 0.21 | 0.61 | 1.8 | 0.1 | 0.71 | 0.18 | 0.06 | 0.75 | 0.97 | 1.23 | 0.01 | 0.67 | 1.4 | 1.01 | 1.5 | 0.55 | 4.49
| Analytical results for Al in OREAS 313 (Certified Value 4.9 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 4.61 | 4.83 | 5.06 | 4.4 | 4.8118 | 4.8584 | 5.1 | 4.83 | 4.81 | 4.95 | 4.67 | 4.83 | 4.79 | 5.0261 | 5.01 | 4.7277 | 5.15 | 4.6379 | 5.15 | 5.03 | 4.95 | 4.78 | 5.0833 | 4.88 | 5.36 | 1
| 1 | 2 | 4.69 | 4.8208 | 5 | 4.46 | 4.5193 | 4.8971 | 5.13 | 4.9 | 4.82 | 4.9 | 4.68 | 4.79 | 4.82 | 5.021 | 5.13 | 4.7759 | 5.26 | 4.6672 | 5.04 | 5.08 | 5.09 | 4.84 | 5.0841 | 4.85 | 5.38 | #N/A
| 1 | 3 | 4.66 | 4.8485 | 5.16 | 4.55 | 4.6447 | 4.8362 | 5.1 | 5.05 | 4.84 | 4.98 | 4.74 | 4.81 | 4.91 | 5.0159 | 5.09 | 4.7516 | 5.26 | 4.6097 | 5.06 | 5.03 | 4.9 | 4.82 | 5.1131 | 4.8 | 5.35 | 16
| 1 | 4 | 4.7 | 4.8366 | 4.935 | 4.45 | 4.8798 | 4.9658 | 5.05 | 4.85 | 4.83 | 5.01 | 4.71 | 4.75 | 4.75 | 5.0333 | 4.9 | 4.6848 | 5.13 | 4.5604 | 5.17 | 4.88 | 4.91 | 4.81 | 5.026 | 5.08 | 5.46 | 4.8992
| 1 | 5 | 4.66 | 4.83 | 4.98 | 4.49 | 4.5949 | 4.8092 | 5.19 | 4.71 | 4.89 | 4.93 | 4.76 | 4.75 | 4.83 | 5.0284 | 4.93 | 4.7695 | 5.26 | 4.6726 | 4.75 | 5.17 | 4.98 | 4.79 | 5.0317 | 4.86 | 5.46 | 8
| 1 | 6 | 4.69 | 4.8202 | 4.9 | 4.47 | 4.681 | 4.9569 | 5.13 | 4.96 | 4.85 | 4.89 | 4.76 | 4.75 | 4.8 | 5.02 | 4.98 | 4.7603 | 5.3 | 4.5401 | 5.12 | 5.02 | 4.95 | 4.88 | 5.0758 | 4.82 | 5.23
| Mean |  | 4.6683 | 4.831 | 5.0058 | 4.47 | 4.6886 | 4.8873 | 5.1167 | 4.8833 | 4.84 | 4.9433 | 4.72 | 4.78 | 4.8167 | 5.0241 | 5.0067 | 4.745 | 5.2267 | 4.6147 | 5.0483 | 5.035 | 4.9633 | 4.82 | 5.069 | 4.8817 | 5.3733
| Median |  | 4.675 | 4.83 | 4.99 | 4.465 | 4.6629 | 4.8778 | 5.115 | 4.875 | 4.835 | 4.94 | 4.725 | 4.77 | 4.81 | 5.0236 | 4.995 | 4.756 | 5.26 | 4.6238 | 5.09 | 5.03 | 4.95 | 4.815 | 5.0796 | 4.855 | 5.37
| Std Dev. |  | 0.033116 | 0.010586 | 0.093403 | 0.049396 | 0.13503 | 0.064261 | 0.046332 | 0.11656 | 0.028284 | 0.046332 | 0.039497 | 0.035214 | 0.053541 | 0.006337 | 0.089592 | 0.033929 | 0.069186 | 0.05511 | 0.15459 | 0.094393 | 0.068605 | 0.036332 | 0.033657 | 0.10128 | 0.085245
| Rel.Std.Dev. |  | 0.7094% | 0.2191% | 1.866% | 1.105% | 2.88% | 1.315% | 0.9055% | 2.387% | 0.5844% | 0.9373% | 0.8368% | 0.7367% | 1.112% | 0.1261% | 1.789% | 0.7151% | 1.324% | 1.194% | 3.062% | 1.875% | 1.382% | 0.7538% | 0.664% | 2.075% | 1.586%
| PDM3 |  | -4.711% | -1.391% | 2.178% | -8.76% | -4.298% | -0.2423% | 4.44% | -0.3229% | -1.207% | 0.9018% | -3.657% | -2.432% | -1.684% | 2.551% | 2.195% | -3.147% | 6.685% | -5.807% | 3.045% | 2.773% | 1.31% | -1.616% | 3.467% | -0.3569% | 9.679%
| Z-Score (Absolute) |  | 1.05 | 0.25 | 0.6 | 2.01 | 0.95 | 0.02 | 1.14 | 0 | 0.21 | 0.29 | 0.8 | 0.5 | 0.32 | 0.69 | 0.6 | 0.67 | 1.67 | 1.31 | 0.8 | 0.74 | 0.39 | 0.31 | 0.9 | 0.01 | 2.39
| Analytical results for As in OREAS 313 (Certified Value 94 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 87 | 96.57 | 95.46 | 63 | 94.5 | 92.4151 | 90.5 | 88.7 | 93.9 | 98.4 | 90.9 | 111 | 85 | 93.1718 | 109 | 96.3 | 92 | 91.7 | 99 | 93 | 95 | 95.1 | 97.4 | 96.9 | 42.4 | 1
| 1 | 2 | 89 | 95.39 | 93.6 | 59.92 | 90.6 | 93.7497 | 89.2 | 87.9 | 93 | 105 | 94.3 | 94.7 | 85 | 96.3296 | 104 | 98.3 | 92.8 | 93.4 | 98 | 92 | 95 | 92.7 | 96.1 | 96.7 | 43.8 | 11
| 1 | 3 | 90 | 89.53 | 90 | 63.26 | 91.9 | 91.9746 | 89.5 | 90.5 | 91.8 | 98 | 93.7 | 100 | 87 | 95.763 | 113 | 99.3 | 94 | 85 | 97 | 93 | 94 | 98.1 | 100.8 | 95.7 | 44.4 | 16
| 1 | 4 | 89 | 98.83 | 94 | 58.34 | 92.1 | 92.594 | 87.7 | 89.9 | 92.9 | 99.2 | 92.6 | 96 | 87 | 94.9199 | 107 | 95.2 | 90.9 | 89 | 99 | 94 | 95 | 94.9 | 96.3 | 96.2 | 44.1 | 93.8022
| 1 | 5 | 89 | 95.07 | 97.2 | 59.92 | 91.8 | 90.3207 | 91.5 | 94.1 | 96.2 | 95 | 95.9 | 100 | 88 | 97.093 | 116 | 98.7 | 93.7 | 93.8 | 92 | 94 | 96 | 93.5 | 97.3 | 95.5 | 41.6 | 9
| 1 | 6 | 89 | 97.04 | 94.5 | 64.44 | 93.9 | 93.134 | 93.2 | 89.7 | 96.4 | 95.6 | 94.3 | 91.3 | 88 | 93.1669 | 106 | 98.4 | 93.4 | 88.1 | 100 | 93 | 96 | 94.4 | 97.3 | 95.5 | 43.6
| Mean |  | 88.8333 | 95.405 | 94.1267 | 61.48 | 92.4667 | 92.3647 | 90.2667 | 90.1333 | 94.0333 | 98.5333 | 93.6167 | 98.8333 | 86.6667 | 95.074 | 109.1667 | 97.7 | 92.8 | 90.1667 | 97.5 | 93.1667 | 95.1667 | 94.7833 | 97.5333 | 96.0833 | 43.3167
| Median |  | 89 | 95.98 | 94.25 | 61.46 | 92 | 92.5046 | 90 | 89.8 | 93.45 | 98.2 | 94 | 98 | 87 | 95.3415 | 108 | 98.35 | 93.1 | 90.35 | 98.5 | 93 | 95 | 94.65 | 97.3 | 95.95 | 43.7
| Std Dev. |  | 0.98319 | 3.1746 | 2.3952 | 2.407 | 1.4542 | 1.1743 | 1.9232 | 2.1519 | 1.879 | 3.568 | 1.7069 | 6.8208 | 1.3663 | 1.6375 | 4.535 | 1.5887 | 1.1713 | 3.4157 | 2.881 | 0.75277 | 0.75277 | 1.8573 | 1.6955 | 0.61455 | 1.0852
| Rel.Std.Dev. |  | 1.107% | 3.327% | 2.545% | 3.915% | 1.573% | 1.271% | 2.131% | 2.387% | 1.998% | 3.621% | 1.823% | 6.901% | 1.576% | 1.722% | 4.154% | 1.626% | 1.262% | 3.788% | 2.955% | 0.808% | 0.791% | 1.96% | 1.738% | 0.6396% | 2.505%
| PDM3 |  | -5.297% | 1.709% | 0.3459% | -34.46% | -1.424% | -1.533% | -3.769% | -3.911% | 0.2464% | 5.044% | -0.1978% | 5.364% | -7.607% | 1.356% | 16.38% | 4.155% | -1.068% | -3.876% | 3.942% | -0.6775% | 1.455% | 1.046% | 3.978% | 2.432% | -53.82%
| Z-Score (Absolute) |  | 1.01 | 0.27 | 0.02 | 6.33 | 0.3 | 0.32 | 0.73 | 0.76 | 0 | 0.88 | 0.08 | 0.93 | 1.43 | 0.2 | 2.94 | 0.71 | 0.24 | 0.75 | 0.67 | 0.17 | 0.22 | 0.15 | 0.68 | 0.4 | 9.87
| Analytical results for B in OREAS 313 (Indicative Value < 10 ppm) | [26 empty] | Indicative
B | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.20 | [24 empty] | ppm
|  |  | 4A*MS | [24 empty] | 1
1
| 1 | 1 | < 10 | [24 empty] | 1
| 1 | 2 | < 10 | [24 empty] | 1
| 1 | 3 | < 10 | [24 empty] | 16
| 1 | 4 | < 10 | [24 empty] | < 10
| 1 | 5 | < 10 | [24 empty] | 7
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 313 (Indicative Value 353 ppm) | [26 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
0
| 1 | 1 | 918 | 2356 | 1934.27 | 253.2 | 180 | 100 | 160 | 770 | 730 | 263 | 92 | 330.131 | 980 | 427.8 | 94 | 23 | 410 | 470 | 1102 | > 1000 | [5 empty] | 1
| 1 | 2 | 910 | 2299 | 1931.68 | 204.3 | 280 | 90 | 170 | 60 | 330 | 131 | 86 | 312.781 | 650 | 352.1 | 72 | 22 | 429 | 477 | 1276 | > 1000 | [5 empty] | 2
| 1 | 3 | 953 | 2375 | 1924.69 | 235.3 | 170 | 80 | 120 | 670 | 220 | 19 | 102 | 314.2063 | 1270 | 455.9 | 79 | 21 | 404 | 474 | 1166 | > 1000 | [5 empty] | 16
| 1 | 4 | 935 | 2300 | 1933.93 | 271.9 | 1490 | 90 | 50 | 700 | 300 | 71 | 104 | 315.5279 | 1190 | 363.3 | 63 | 23 | 403 | 476 | 1250 | > 1000 | [5 empty] | 353.1194
| 1 | 5 | 914 | 2299 | 1925.41 | 262.8 | 230 | 60 | 120 | 950 | 240 | 31 | 111 | 346.4498 | 950 | 380.4 | 93 | 26 | 423 | 476 | 1123 | > 1000 | [5 empty] | 8
| 1 | 6 | 950 | 2280 | 1933.93 | 235.5 | 230 | 110 | 230 | 660 | 310 | 309 | 79 | 327.1669 | 1230 | 497.7 | 65 | 19 | 381 | 474 | 1288 | > 1000
| Mean |  | 930 | 2318.1667 | 1930.6517 | 243.8333 | 430 | 88.3333 | 141.6667 | 635 | 355 | 137.3333 | 95.6667 | 324.3771 | 1045 | 412.8667 | 77.6667 | 22.3333 | 408.3333 | 474.5 | 1200.8333
| Median |  | 926.5 | 2299.5 | 1932.805 | 244.35 | 230 | 90 | 140 | 685 | 305 | 101 | 97 | 321.3474 | 1085 | 404.1 | 75.5 | 22.5 | 407 | 475 | 1208
| Std Dev. |  | 18.7297 | 37.9021 | 4.4421 | 24.2502 | 520.8071 | 17.224 | 60.4704 | 301.3802 | 188.5471 | 122.4837 | 12.0775 | 12.9947 | 234.5847 | 57.4023 | 13.5006 | 2.3381 | 16.9666 | 2.51 | 80.8762
| Rel.Std.Dev. |  | 2.014% | 1.635% | 0.2301% | 9.945% | 121.1% | 19.5% | 42.68% | 47.46% | 53.11% | 89.19% | 12.62% | 4.006% | 22.45% | 13.9% | 17.38% | 10.47% | 4.155% | 0.529% | 6.735%
| PDM3 |  | 163.4% | 556.5% | 446.7% | -30.95% | 21.77% | -74.98% | -59.88% | 79.83% | 0.5326% | -61.11% | -72.91% | -8.14% | 195.9% | 16.92% | -78.01% | -93.68% | 15.64% | 34.37% | 240.1%
| Z-Score (Absolute) |  | 1.3 | 4.75 | 3.79 | 0.41 | 0.05 | 0.8 | 0.66 | 0.56 | 0.13 | 0.67 | 0.78 | 0.21 | 1.58 | 0.01 | 0.82 | 0.96 | 0 | 0.16 | 1.97 | NA
| Analytical results for Be in OREAS 313 (Certified Value 2.19 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2 | 1.8991 | 2 | 2.38 | 2.24 | 2.3515 | 2.24 | 2.15 | 2.15 | 2.17 | 2.05 | 2.5 | 1.95 | 2.1995 | 1.48 | 2.18 | 2 | 3 | 1.8 | 2 | 2.22 | 2.21 | 2.41 | 2 |  | 1
| 1 | 2 | 2 | 1.798 | 2 | 2.32 | 2.19 | 2.3047 | 2.25 | 2.19 | 2.15 | 2.13 | 2.08 | 2.3 | 2.05 | 2.1325 | 1.47 | 2.42 | 2 | 2 | 1.9 | 2 | 2.14 | 2.33 | 2.45 | 2 |  | 33
| 1 | 3 | 2 | 1.8074 | 2 | 2.43 | 2.18 | 2.3057 | 2.25 | 2.24 | 2.17 | 2.17 | 2.1 | 2.4 | 2.06 | 2.1984 | 1.37 | 2.29 | 2 | 2 | 1.8 | 2 | 2.18 | 2.22 | 2.4 | 2 |  | 16
| 1 | 4 | 2 | 1.8012 | 2 | 2.47 | 2.41 | 2.4362 | 2.21 | 2.17 | 2.15 | 2.23 | 2.09 | 2.2 | 1.96 | 2.1428 | 1.37 | 2.49 | 2 | 3 | 1.8 | 2 | 2.23 | 2.22 | 2.38 | 2 |  | 2.1866
| 1 | 5 | 2 | 1.8679 | 2 | 2.38 | 2.27 | 2.3234 | 2.26 | 2.2 | 2.17 | 2.17 | 2.12 | 2.2 | 2.06 | 2.1433 | 1.21 | 2.34 | 2 | 2 | 1.8 | 2 | 2.25 | 2.21 | 2.28 | 2 |  | 10
| 1 | 6 | 2 | 1.8237 | 2 | 2.4 | 2.18 | 2.3733 | 2.24 | 2.21 | 2.17 | 2.17 | 2.1 | 2.3 | 2.11 | 2.1222 | 1.54 | 2.46 | 2 | 3 | 1.8 | 2 | 2.13 | 2.24 | 2.33 | 2
| Mean |  | 2 | 1.8329 | 2 | 2.3967 | 2.245 | 2.3491 | 2.2417 | 2.1933 | 2.16 | 2.1733 | 2.09 | 2.3167 | 2.0317 | 2.1564 | 1.4067 | 2.3633 | 2 | 2.5 | 1.8167 | 2 | 2.1917 | 2.2383 | 2.375 | 2
| Median |  | 2 | 1.8155 | 2 | 2.39 | 2.215 | 2.3375 | 2.245 | 2.195 | 2.16 | 2.17 | 2.095 | 2.3 | 2.055 | 2.143 | 1.42 | 2.38 | 2 | 2.5 | 1.8 | 2 | 2.2 | 2.22 | 2.39 | 2
| Std Dev. |  | 0 | 0.041398 | 0 | 0.050859 | 0.088713 | 0.050363 | 0.017224 | 0.031411 | 0.010954 | 0.032042 | 0.023664 | 0.1169 | 0.063061 | 0.03382 | 0.11708 | 0.11673 | 0 | 0.54772 | 0.040825 | 0 | 0.049565 | 0.046224 | 0.06091 | 0
| Rel.Std.Dev. |  | 0% | 2.259% | 0% | 2.122% | 3.952% | 2.144% | 0.7684% | 1.432% | 0.5072% | 1.474% | 1.132% | 5.046% | 3.104% | 1.568% | 8.323% | 4.939% | 0% | 21.91% | 2.247% | 0% | 2.262% | 2.065% | 2.565% | 0%
| PDM3 |  | -8.533% | -16.18% | -8.533% | 9.608% | 2.672% | 7.434% | 2.52% | 0.309% | -1.215% | -0.6056% | -4.417% | 5.95% | -7.085% | -1.378% | -35.67% | 8.084% | -8.533% | 14.33% | -16.92% | -8.533% | 0.2328% | 2.367% | 8.617% | -8.533%
| Z-Score (Absolute) |  | NA | 2.14 | NA | 1.21 | 0.31 | 0.93 | 0.29 | 0 | 0.19 | 0.11 | 0.61 | 0.74 | 0.96 | 0.21 | 4.68 | 1.02 | NA | NA | 2.24 | NA | 0 | 0.27 | 1.09 | NA
Results from laboratories 2.01, 2.03, 2.17, 2.19, 2.21 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 313 (Certified Value 0.43 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 0.4 | 0.4249 | 0.6 | < 0.5 | 0.44 | 0.43114 | 0.45 | 0.44 | 0.43 | 0.44 | 0.43 | 0.44 | 0.4 | 0.4222 | 0.64 | 0.43 | 0.41 | 0.49 | 0.4 | 0.9 | 0.44 | 0.42 | 0.46 | 0.44 | 0.6 | 1
| 1 | 2 | 0.4 | 0.4416 | 0.5 | < 0.5 | 0.44 | 0.43932 | 0.44 | 0.46 | 0.43 | 0.48 | 0.43 | 0.43 | 0.43 | 0.4499 | 0.48 | 0.44 | 0.4 | 0.52 | 0.5 | 0.93 | 0.43 | 0.38 | 0.48 | 0.44 | 0.6 | 34
| 1 | 3 | 0.41 | 0.4461 | 0.5 | < 0.5 | 0.44 | 0.45443 | 0.46 | 0.44 | 0.43 | 0.45 | 0.43 | 0.4 | 0.42 | 0.4429 | 1.17 | 0.44 | 0.39 | 0.52 | 0.5 | 0.95 | 0.43 | 0.41 | 0.49 | 0.43 | 0.6 | 16
| 1 | 4 | 0.4 | 0.3973 | 0.5 | < 0.5 | 0.46 | 0.45952 | 0.45 | 0.56 | 0.43 | 0.48 | 0.41 | 0.4 | 0.41 | 0.4047 | 0.48 | 0.45 | 0.4 | 0.48 | 0.5 | 0.88 | 0.42 | 0.4 | 0.47 | 0.44 | 0.6 | 0.43355
| 1 | 5 | 0.4 | 0.4377 | 0.5 | < 0.5 | 0.46 | 0.44087 | 0.47 | 0.44 | 0.43 | 0.44 | 0.44 | 0.43 | 0.42 | 0.4347 | 0.48 | 0.46 | 0.4 | 0.51 | 0.4 | 0.95 | 0.43 | 0.41 | 0.49 | 0.45 | 0.5 | 11
| 1 | 6 | 0.41 | 0.4612 | 0.5 | < 0.5 | 0.46 | 0.41837 | 0.45 | 0.41 | 0.44 | 0.43 | 0.43 | 0.42 | 0.44 | 0.4181 | 0.51 | 0.42 | 0.41 | 0.48 | 0.5 | 0.92 | 0.42 | 0.47 | 0.48 | 0.43 | 0.6
| Mean |  | 0.40333 | 0.4348 | 0.51667 |  | 0.45 | 0.44061 | 0.45333 | 0.45833 | 0.43167 | 0.45333 | 0.42833 | 0.42 | 0.42 | 0.42875 | 0.62667 | 0.44 | 0.40167 | 0.5 | 0.46667 | 0.92167 | 0.42833 | 0.415 | 0.47833 | 0.43833 | 0.58333
| Median |  | 0.4 | 0.43965 | 0.5 |  | 0.45 | 0.4401 | 0.45 | 0.44 | 0.43 | 0.445 | 0.43 | 0.425 | 0.42 | 0.42845 | 0.495 | 0.44 | 0.4 | 0.5 | 0.5 | 0.925 | 0.43 | 0.41 | 0.48 | 0.44 | 0.6
| Std Dev. |  | 0.005164 | 0.021834 | 0.040825 |  | 0.010954 | 0.015064 | 0.010328 | 0.052313 | 0.0040825 | 0.021602 | 0.0098319 | 0.016733 | 0.014142 | 0.016838 | 0.27333 | 0.014142 | 0.0075277 | 0.018974 | 0.05164 | 0.027869 | 0.0075277 | 0.030166 | 0.01169 | 0.0075277 | 0.040825
| Rel.Std.Dev. |  | 1.28% | 5.022% | 7.902% |  | 2.434% | 3.419% | 2.278% | 11.41% | 0.9457% | 4.765% | 2.295% | 3.984% | 3.367% | 3.927% | 43.62% | 3.214% | 1.874% | 3.795% | 11.07% | 3.024% | 1.757% | 7.269% | 2.444% | 1.717% | 6.999%
| PDM3 |  | -6.969% | 0.2893% | 19.17% |  | 3.795% | 1.629% | 4.564% | 5.717% | -0.4335% | 4.564% | -1.202% | -3.124% | -3.124% | -1.106% | 44.54% | 1.489% | -7.353% | 15.33% | 7.639% | 112.6% | -1.202% | -4.278% | 10.33% | 1.104% | 34.55%
| Z-Score (Absolute) |  | 1.34 | 0.07 | NA | 7.55 | 0.54 | 0.16 | 0.68 | 0.88 | 0.2 | 0.68 | 0.33 | 0.67 | 0.67 | 0.32 | 7.69 | 0.14 | 1.41 | 2.57 | NA | 19.63 | 0.33 | 0.87 | 1.69 | 0.07 | NA
Results from laboratories 2.03, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 313 (Certified Value 1.35 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.3 | 1.3447 | 1.388 | 1.33 | 1.3367 | 1.3532 | 1.46 | 1.36 | 1.34 | 1.39 | 1.31 | 1.45 | 1.374 | 1.3387 | 1.47 | 1.3447 | 1.41 | 1.353 | 1.37 | 1.37 | 1.33 | 1.33 | 1.4403 | 1.33 | 1.47 | 1
| 1 | 2 | 1.32 | 1.3105 | 1.37 | 1.31 | 1.2559 | 1.3626 | 1.46 | 1.38 | 1.34 | 1.37 | 1.33 | 1.36 | 1.369 | 1.3281 | 1.5 | 1.3843 | 1.43 | 1.3534 | 1.34 | 1.37 | 1.33 | 1.35 | 1.4311 | 1.34 | 1.49 | #N/A
| 1 | 3 | 1.3 | 1.3509 | 1.41 | 1.33 | 1.2616 | 1.34 | 1.46 | 1.42 | 1.36 | 1.39 | 1.34 | 1.31 | 1.409 | 1.3193 | 1.5 | 1.4092 | 1.43 | 1.3302 | 1.35 | 1.37 | 1.32 | 1.34 | 1.4305 | 1.31 | 1.49 | 16
| 1 | 4 | 1.32 | 1.3464 | 1.3503 | 1.33 | 1.3726 | 1.362 | 1.44 | 1.35 | 1.37 | 1.43 | 1.33 | 1.27 | 1.362 | 1.3384 | 1.43 | 1.3676 | 1.41 | 1.3222 | 1.37 | 1.32 | 1.33 | 1.33 | 1.4255 | 1.39 | 1.48 | 1.3547
| 1 | 5 | 1.31 | 1.3268 | 1.378 | 1.33 | 1.2707 | 1.3538 | 1.49 | 1.33 | 1.36 | 1.39 | 1.35 | 1.29 | 1.393 | 1.3475 | 1.46 | 1.4083 | 1.43 | 1.3507 | 1.26 | 1.38 | 1.32 | 1.32 | 1.4382 | 1.32 | 1.5 | 12
| 1 | 6 | 1.32 | 1.3374 | 1.346 | 1.32 | 1.3141 | 1.385 | 1.47 | 1.4 | 1.36 | 1.37 | 1.33 | 1.35 | 1.372 | 1.3189 | 1.47 | 1.3918 | 1.44 | 1.3248 | 1.36 | 1.35 | 1.33 | 1.35 | 1.4202 | 1.32 | 1.48
| Mean |  | 1.3117 | 1.3361 | 1.3737 | 1.325 | 1.3019 | 1.3594 | 1.4633 | 1.3733 | 1.355 | 1.39 | 1.3317 | 1.3383 | 1.3798 | 1.3318 | 1.4717 | 1.3843 | 1.425 | 1.3391 | 1.3417 | 1.36 | 1.3267 | 1.3367 | 1.431 | 1.335 | 1.485
| Median |  | 1.315 | 1.3411 | 1.374 | 1.33 | 1.2924 | 1.3579 | 1.46 | 1.37 | 1.36 | 1.39 | 1.33 | 1.33 | 1.373 | 1.3333 | 1.47 | 1.388 | 1.43 | 1.3405 | 1.355 | 1.37 | 1.33 | 1.335 | 1.4308 | 1.325 | 1.485
| Std Dev. |  | 0.0098319 | 0.015111 | 0.023955 | 0.0083666 | 0.047057 | 0.014986 | 0.01633 | 0.033267 | 0.012247 | 0.021909 | 0.013292 | 0.064627 | 0.017634 | 0.01161 | 0.026394 | 0.024897 | 0.012247 | 0.014843 | 0.041673 | 0.021909 | 0.005164 | 0.012111 | 0.0075561 | 0.02881 | 0.010488
| Rel.Std.Dev. |  | 0.7496% | 1.131% | 1.744% | 0.6314% | 3.614% | 1.102% | 1.116% | 2.422% | 0.9039% | 1.576% | 0.9981% | 4.829% | 1.278% | 0.8718% | 1.794% | 1.799% | 0.8595% | 1.108% | 3.106% | 1.611% | 0.3892% | 0.906% | 0.528% | 2.158% | 0.7063%
| PDM3 |  | -3.177% | -1.373% | 1.403% | -2.193% | -3.895% | 0.3501% | 8.019% | 1.375% | 0.02172% | 2.605% | -1.701% | -1.209% | 1.855% | -1.69% | 8.634% | 2.186% | 5.189% | -1.156% | -0.9625% | 0.3908% | -2.07% | -1.332% | 5.629% | -1.455% | 9.618%
| Z-Score (Absolute) |  | 1.25 | 0.55 | 0.54 | 0.87 | 1.53 | 0.13 | 3.13 | 0.53 | 0 | 1.01 | 0.67 | 0.48 | 0.72 | 0.67 | 3.37 | 0.85 | 2.02 | 0.46 | 0.39 | 0.14 | 0.82 | 0.53 | 2.2 | 0.58 | 3.76
| Analytical results for Cd in OREAS 313 (Certified Value 113 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 107 | 88.7025 | 103.6 | 110.67 | 107.59 | 117.6985 | 124.5 | 108.5 | 110 | 111.5 | 108.5 | 118 | 112 | 113.7044 | 112 | 113.48 | 115.98 | 105.91 | 116.7 | 118 | 87.9 | 116.65 | 121.58 | 79.6 |  | 1
| 1 | 2 | 106 | 85.935 | 102 | 110.48 | 104.66 | 118.7942 | 123 | 111 | 109.5 | 113.5 | 111.5 | 118 | 110 | 112.9611 | 113 | 118.47 | 116.38 | 110.98 | 114.5 | 119 | 91 | 115.4 | 124.02 | 82.1 |  | 13
| 1 | 3 | 107 | 88.2225 | 101.8 | 111.54 | 106.84 | 115.3198 | 124 | 116.5 | 110 | 113 | 112 | 117 | 113 | 113.5565 | 113 | 118.42 | 115 | 99.12 | 114.2 | 118 | 88.2 | 117.78 | 120.56 | 82.7 |  | 16
| 1 | 4 | 107 | 83.22 | 102 | 110.52 | 111.44 | 117.6302 | 122 | 112 | 108 | 115 | 110.5 | 117 | 112 | 112.8921 | 109 | 115.84 | 111.27 | 101.11 | 118 | 116 | 88.8 | 115.11 | 122.28 | 83 |  | 113.1288
| 1 | 5 | 107 | 88.11 | 103 | 109.04 | 107.54 | 115.2974 | 125 | 107 | 111.5 | 110.5 | 113.5 | 116 | 114 | 113.9884 | 109 | 118.11 | 114.5 | 106.72 | 109.6 | 118 | 88.6 | 117.94 | 123.8 | 78.6 |  | 13
| 1 | 6 | 106 | 87.165 | 103 | 108.52 | 108.35 | 118.0783 | 124.5 | 110.5 | 111.5 | 109.5 | 113.5 | 116 | 114 | 113.4469 | 110 | 117.08 | 115.07 | 104.76 | 119.5 | 118 | 90.2 | 114.55 | 122.28 | 79.2
| Mean |  | 106.6667 | 86.8925 | 102.5667 | 110.1283 | 107.7367 | 117.1364 | 123.8333 | 110.9167 | 110.0833 | 112.1667 | 111.5833 | 117 | 112.5 | 113.4249 | 111 | 116.9 | 114.7 | 104.7667 | 115.4167 | 117.8333 | 89.1167 | 116.2383 | 122.42 | 80.8667
| Median |  | 107 | 87.6375 | 102.5 | 110.5 | 107.565 | 117.6643 | 124.25 | 110.75 | 110 | 112.25 | 111.75 | 117 | 112.5 | 113.5017 | 111 | 117.595 | 115.035 | 105.335 | 115.6 | 118 | 88.7 | 116.025 | 122.28 | 80.85
| Std Dev. |  | 0.5164 | 2.0497 | 0.73121 | 1.1256 | 2.21 | 1.4749 | 1.1255 | 3.2774 | 1.3197 | 2.0412 | 1.9083 | 0.89443 | 1.5166 | 0.42719 | 1.8974 | 1.9541 | 1.8162 | 4.2186 | 3.4971 | 0.98319 | 1.2172 | 1.4331 | 1.3167 | 1.947
| Rel.Std.Dev. |  | 0.4841% | 2.359% | 0.7129% | 1.022% | 2.051% | 1.259% | 0.9089% | 2.955% | 1.199% | 1.82% | 1.71% | 0.7645% | 1.348% | 0.3766% | 1.709% | 1.672% | 1.583% | 4.027% | 3.03% | 0.8344% | 1.366% | 1.233% | 1.076% | 2.408%
| PDM3 |  | -5.712% | -23.19% | -9.336% | -2.652% | -4.766% | 3.543% | 9.462% | -1.955% | -2.692% | -0.8504% | -1.366% | 3.422% | -0.5558% | 0.2618% | -1.882% | 3.334% | 1.389% | -7.392% | 2.022% | 4.159% | -21.23% | 2.749% | 8.213% | -28.52%
| Z-Score (Absolute) |  | 0.75 | 3.59 | 1.34 | 0.25 | 0.59 | 0.76 | 1.72 | 0.14 | 0.26 | 0.04 | 0.04 | 0.74 | 0.09 | 0.22 | 0.13 | 0.72 | 0.41 | 1.02 | 0.51 | 0.86 | 3.27 | 0.63 | 1.51 | 4.45
| Analytical results for Ce in OREAS 313 (Certified Value 56 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 0
1
| 1 | 1 | 55.5418 | 60 | 54.9 | 59.34 | 60.4649 | 59.8 | 55.3 | 57.9 | 55.7 | 50.1 | 59.8 | 49.6 | 36.1 | 58.44 | 54.74 | 57 | 45.61 | 61.4 | 56.6 | 62.36 | 63.45 | 46 |  |  |  | 1
| 1 | 2 | 55.4245 | 62 | 56.8 | 58.31 | 62.5198 | 60.5 | 53.5 | 57 | 57.7 | 51.7 | 52 | 48.34 | 37.5 | 58.13 | 54.97 | 55 | 44.87 | 60.1 | 55.3 | 61.7 | 62.86 | 49 |  |  |  | 14
| 1 | 3 | 55.5576 | 60 | 59.7 | 60.19 | 59.7424 | 58.7 | 55.2 | 55.2 | 54.5 | 53.4 | 53.6 | 51.07 | 37.3 | 59.58 | 54.57 | 57 | 43.88 | 59.8 | 57.8 | 61.57 | 63.09 | 45 |  |  |  | 16
| 1 | 4 | 55.4507 | 60 | 58.9 | 60.93 | 61.2501 | 59.3 | 54.7 | 57.3 | 56 | 51.8 | 52 | 49.7 | 35.9 | 56.43 | 52.53 | 57 | 45.88 | 59.8 | 56.3 | 61.38 | 63.14 | 49 |  |  |  | 56.1653
| 1 | 5 | 55.5434 | 59 | 57.8 | 60.31 | 59.5555 | 60.1 | 57.1 | 57.9 | 53.2 | 54 | 56 | 52.79 | 37.5 | 59.68 | 53.72 | 53 | 44.97 | 64.5 | 54.7 | 60.71 | 62.54 | 42 |  |  |  | 14
| 1 | 6 | 55.5732 | 61 | 57.1 | 60.99 | 60.8341 | 61.7 | 54.9 | 58.7 | 52.6 | 51.6 | 50.4 | 50.49 | 36.6 | 58.67 | 51.6 | 58 | 44.39 | 61.2 | 55.8 | 61.68 | 60.58 | 48
| Mean |  | 55.5152 | 60.3333 | 57.5333 | 60.0117 | 60.7278 | 60.0167 | 55.1167 | 57.3333 | 54.95 | 52.1 | 53.9667 | 50.3317 | 36.8167 | 58.4883 | 53.6883 | 56.1667 | 44.9333 | 61.1333 | 56.0833 | 61.5667 | 62.61 | 46.5
| Median |  | 55.5426 | 60 | 57.45 | 60.25 | 60.6495 | 59.95 | 55.05 | 57.6 | 55.1 | 51.75 | 52.8 | 50.095 | 36.95 | 58.555 | 54.145 | 57 | 44.92 | 60.65 | 56.05 | 61.625 | 62.975 | 47
| Std Dev. |  | 0.061727 | 1.0328 | 1.6908 | 1.0265 | 1.0872 | 1.0362 | 1.1669 | 1.1978 | 1.8982 | 1.4 | 3.4279 | 1.5175 | 0.71671 | 1.1843 | 1.3587 | 1.8348 | 0.74352 | 1.7907 | 1.0834 | 0.53455 | 1.0397 | 2.7386
| Rel.Std.Dev. |  | 0.1112% | 1.712% | 2.939% | 1.71% | 1.79% | 1.726% | 2.117% | 2.089% | 3.454% | 2.687% | 6.352% | 3.015% | 1.947% | 2.025% | 2.531% | 3.267% | 1.655% | 2.929% | 1.932% | 0.8683% | 1.661% | 5.889%
| PDM3 |  | -1.158% | 7.421% | 2.436% | 6.848% | 8.123% | 6.857% | -1.867% | 2.08% | -2.164% | -7.238% | -3.915% | -10.39% | -34.45% | 4.136% | -4.41% | 0.002347% | -20% | 8.845% | -0.146% | 9.617% | 11.47% | -17.21%
| Z-Score (Absolute) |  | 0.11 | 0.73 | 0.24 | 0.67 | 0.8 | 0.67 | 0.17 | 0.21 | 0.2 | 0.7 | 0.37 | 1 | 3.35 | 0.41 | 0.42 | 0.01 | 1.94 | 0.87 | 0.01 | 0.94 | 1.12 | 1.67
| Analytical results for Co in OREAS 313 (Certified Value 25.5 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 23 | 25.537 | 26 | 25.42 | 25.9 | 26.1632 | 25.7 | 23.5 | 25.8 | 25 | 23.3 | 27 | 22.1 | 28.9085 | 22.7 | 25.7 | 25.8 | 23.6 | 27.1 | 20.4 | 26.5 | 25.32 | 25.9 | 28.9 | 24.4 | 1
| 1 | 2 | 24 | 25.4624 | 26.6 | 25.2 | 24.8 | 26.7604 | 25.4 | 23.1 | 26.1 | 26.3 | 24.2 | 27 | 22.7 | 28.0046 | 25.2 | 26 | 26.4 | 24 | 27.1 | 21.1 | 27.3 | 24.85 | 25.8 | 29.3 | 25 | 16
| 1 | 3 | 24 | 25.5394 | 25.6 | 25.93 | 24.9 | 25.4191 | 25.5 | 23.7 | 25.7 | 25.1 | 24.8 | 27 | 23.2 | 28.7652 | 25 | 26.2 | 25.7 | 22.1 | 27.2 | 20.4 | 26.5 | 25.88 | 26.4 | 29.1 | 25.1 | 16
| 1 | 4 | 24 | 25.5191 | 26.2 | 25.03 | 26.8 | 26.2641 | 25.1 | 23.6 | 25.3 | 25.9 | 23.7 | 27 | 23.1 | 28.3743 | 25.2 | 25.9 | 25.4 | 22.6 | 27.6 | 20.2 | 26.4 | 24.96 | 25.8 | 29.6 | 25.3 | 25.5278
| 1 | 5 | 24 | 25.5132 | 25.4 | 25.22 | 25.2 | 25.4467 | 26.1 | 22.8 | 27.2 | 24.7 | 25.3 | 26 | 23 | 28.5663 | 25.5 | 26.6 | 25.7 | 24.2 | 25.7 | 20.6 | 26.7 | 24.92 | 26.1 | 29.3 | 24 | 15
| 1 | 6 | 24 | 25.4355 | 26.3 | 25.31 | 25.4 | 26.1632 | 26.4 | 23.7 | 26.5 | 24.4 | 24.1 | 27 | 22.9 | 28.1102 | 26.2 | 25.9 | 25.8 | 22.7 | 28.3 | 20.7 | 26.1 | 25.81 | 25.8 | 29.2 | 24.3
| Mean |  | 23.8333 | 25.5011 | 26.0167 | 25.3517 | 25.5 | 26.0361 | 25.7 | 23.4 | 26.1 | 25.2333 | 24.2333 | 26.8333 | 22.8333 | 28.4548 | 24.9667 | 26.05 | 25.8 | 23.2 | 27.1667 | 20.5667 | 26.5833 | 25.29 | 25.9667 | 29.2333 | 24.6833
| Median |  | 24 | 25.5162 | 26.1 | 25.265 | 25.3 | 26.1632 | 25.6 | 23.55 | 25.95 | 25.05 | 24.15 | 27 | 22.95 | 28.4703 | 25.2 | 25.95 | 25.75 | 23.15 | 27.15 | 20.5 | 26.5 | 25.14 | 25.85 | 29.25 | 24.7
| Std Dev. |  | 0.40825 | 0.04249 | 0.44907 | 0.31135 | 0.74833 | 0.51712 | 0.47749 | 0.36878 | 0.67231 | 0.72572 | 0.72572 | 0.40825 | 0.39833 | 0.35848 | 1.1877 | 0.31464 | 0.32863 | 0.85088 | 0.85245 | 0.31411 | 0.40208 | 0.46017 | 0.24221 | 0.23381 | 0.51929
| Rel.Std.Dev. |  | 1.713% | 0.1666% | 1.726% | 1.228% | 2.935% | 1.986% | 1.858% | 1.576% | 2.576% | 2.876% | 2.995% | 1.521% | 1.745% | 1.26% | 4.757% | 1.208% | 1.274% | 3.668% | 3.138% | 1.527% | 1.513% | 1.82% | 0.9328% | 0.7998% | 2.104%
| PDM3 |  | -6.638% | -0.1045% | 1.915% | -0.6899% | -0.1088% | 1.991% | 0.6746% | -8.335% | 2.242% | -1.153% | -5.071% | 5.114% | -10.55% | 11.47% | -2.198% | 2.046% | 1.066% | -9.119% | 6.42% | -19.43% | 4.135% | -0.9314% | 1.719% | 14.52% | -3.308%
| Z-Score (Absolute) |  | 1.88 | 0 | 0.58 | 0.17 | 0 | 0.6 | 0.22 | 2.37 | 0.67 | 0.3 | 1.43 | 1.5 | 3 | 3.33 | 0.6 | 0.62 | 0.34 | 2.59 | 1.88 | 5.56 | 1.22 | 0.24 | 0.52 | 4.2 | 0.92
| Analytical results for Cr in OREAS 313 (Certified Value 69 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 83 | 86.583 | 75 | 75.49 | 67 | 64.3011 | 68 | 64 | 65 | 66 | 69 | 62 | 73 | 75.111 | 59 | 76 | 82 | 57 | 72 | 55 | 49.2 | 76 | 67 | 69 |  | 1
| 1 | 2 | 80 | 84.724 | 77 | 74.7 | 69 | 65.3982 | 76 | 67 | 66 | 70 | 65 | 61 | 73 | 80.2267 | 59 | 78 | 80 | 56 | 72 | 56 | 50.1 | 72 | 67 | 71 |  | 37
| 1 | 3 | 81 | 88.882 | 74 | 77.62 | 70 | 62.4895 | 65 | 68 | 71 | 65 | 67 | 65 | 76 | 72.7477 | 60 | 79 | 79 | 52 | 71 | 55 | 48.4 | 69 | 66 | 68 |  | 16
| 1 | 4 | 81 | 82.755 | 77 | 79.03 | 72 | 65.2429 | 69 | 65 | 69 | 64 | 69 | 74 | 72 | 77.4236 | 59 | 76 | 81 | 57 | 70 | 56 | 47.8 | 74 | 71 | 70 |  | 68.9483
| 1 | 5 | 83 | 87.43 | 75 | 75.5 | 70 | 60.5135 | 67 | 64 | 64 | 67 | 63 | 47 | 72 | 81.0329 | 58 | 79 | 81 | 50 | 73 | 54 | 49 | 71 | 66 | 70 |  | 16
| 1 | 6 | 82 | 84.592 | 75 | 76.03 | 74 | 64.8503 | 70 | 68 | 66 | 66 | 67 | 61 | 72 | 76.1844 | 58 | 77 | 82 | 54 | 70 | 55 | 51 | 68 | 68 | 68
| Mean |  | 81.6667 | 85.8277 | 75.5 | 76.395 | 70.3333 | 63.7993 | 69.1667 | 66 | 66.8333 | 66.3333 | 66.6667 | 61.6667 | 73 | 77.121 | 58.8333 | 77.5 | 80.8333 | 54.3333 | 71.3333 | 55.1667 | 49.25 | 71.6667 | 67.5 | 69.3333
| Median |  | 81.5 | 85.6535 | 75 | 75.765 | 70 | 64.5757 | 68.5 | 66 | 66 | 66 | 67 | 61.5 | 72.5 | 76.804 | 59 | 77.5 | 81 | 55 | 71.5 | 55 | 49.1 | 71.5 | 67 | 69.5
| Std Dev. |  | 1.2111 | 2.2205 | 1.2247 | 1.6168 | 2.4221 | 1.924 | 3.7639 | 1.8974 | 2.6394 | 2.0656 | 2.3381 | 8.7101 | 1.5492 | 3.1331 | 0.75277 | 1.3784 | 1.169 | 2.8752 | 1.2111 | 0.75277 | 1.1554 | 3.0111 | 1.8708 | 1.2111
| Rel.Std.Dev. |  | 1.483% | 2.587% | 1.622% | 2.116% | 3.444% | 3.016% | 5.442% | 2.875% | 3.949% | 3.114% | 3.507% | 14.12% | 2.122% | 4.063% | 1.28% | 1.779% | 1.446% | 5.292% | 1.698% | 1.365% | 2.346% | 4.202% | 2.772% | 1.747%
| PDM3 |  | 18.45% | 24.48% | 9.502% | 10.8% | 2.009% | -7.468% | 0.3167% | -4.276% | -3.067% | -3.793% | -3.309% | -10.56% | 5.876% | 11.85% | -14.67% | 12.4% | 17.24% | -21.2% | 3.459% | -19.99% | -28.57% | 3.943% | -2.101% | 0.5584%
| Z-Score (Absolute) |  | 1.43 | 1.91 | 0.72 | 0.82 | 0.12 | 0.63 | 0.01 | 0.37 | 0.28 | 0.34 | 0.3 | 0.87 | 0.43 | 0.91 | 1.2 | 0.95 | 1.34 | 1.72 | 0.24 | 1.62 | 2.31 | 0.28 | 0.2 | 0.01
| Analytical results for Cs in OREAS 313 (Certified Value 3.12 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 3 | 3.6932 | 3 | 3.14 | 3.2531 | 3.08 | 3.11 | 3.17 | 2.96 | 2.81 | 3.31 | 2.74 | 3.1516 | 3.11 | 3.21 | 3.1 | 2 | 3.13 | 3.19 | 3.18 | 3.2 |  |  |  |  | 1
| 1 | 2 | 3 | 3.5732 | 3.1 | 3.06 | 3.2855 | 3.08 | 3.06 | 3.16 | 3.15 | 2.88 | 3.09 | 2.71 | 3.2312 | 3.1 | 3.21 | 3.2 | 2 | 3.11 | 3.27 | 3.2 | 3.4 |  |  |  |  | 17
| 1 | 3 | 3 | 3.6492 | 3.04 | 3.1 | 3.2138 | 3.1 | 3.16 | 3.1 | 2.99 | 2.93 | 3.11 | 2.77 | 3.259 | 3.18 | 3.22 | 3.1 | 2 | 3.19 | 3.27 | 3.19 | 3.4 |  |  |  |  | 16
| 1 | 4 | 3.1 | 3.584 | 3 | 3.19 | 3.3436 | 2.95 | 3.11 | 3.09 | 3.04 | 2.87 | 3.12 | 2.76 | 3.2872 | 2.96 | 3.18 | 2.9 | 2 | 3.13 | 3.16 | 3.17 | 3.4 |  |  |  |  | 3.1218
| 1 | 5 | 3 | 3.5598 | 3 | 3.11 | 3.1717 | 3.12 | 2.93 | 3.23 | 2.91 | 2.94 | 3.24 | 2.81 | 3.2633 | 3.2 | 3.25 | 3.1 | 2 | 3.04 | 3.36 | 3.2 | 3.1 |  |  |  |  | 17
| 1 | 6 | 3 | 3.6229 | 2.98 | 3.16 | 3.1898 | 3.15 | 3.16 | 3.24 | 2.81 | 2.92 | 3.01 | 2.79 | 3.1433 | 3.04 | 3.15 | 3 | 2 | 3.29 | 3.13 | 3.12 | 3.1
| Mean |  | 3.0167 | 3.6137 | 3.02 | 3.1267 | 3.2429 | 3.08 | 3.0883 | 3.165 | 2.9767 | 2.8917 | 3.1467 | 2.7633 | 3.2226 | 3.0983 | 3.2033 | 3.0667 | 2 | 3.1483 | 3.23 | 3.1767 | 3.2667
| Median |  | 3 | 3.6035 | 3 | 3.125 | 3.2334 | 3.09 | 3.11 | 3.165 | 2.975 | 2.9 | 3.115 | 2.765 | 3.2451 | 3.105 | 3.21 | 3.1 | 2 | 3.13 | 3.23 | 3.185 | 3.3
| Std Dev. |  | 0.040825 | 0.051169 | 0.043818 | 0.046332 | 0.0645 | 0.068993 | 0.08612 | 0.062849 | 0.11553 | 0.048751 | 0.10893 | 0.03559 | 0.060923 | 0.089088 | 0.034448 | 0.10328 | 0 | 0.084479 | 0.085557 | 0.030111 | 0.15055
| Rel.Std.Dev. |  | 1.353% | 1.416% | 1.451% | 1.482% | 1.989% | 2.24% | 2.789% | 1.986% | 3.881% | 1.686% | 3.462% | 1.288% | 1.89% | 2.875% | 1.075% | 3.368% | 0% | 2.683% | 2.649% | 0.9479% | 4.609%
| PDM3 |  | -3.369% | 15.76% | -3.262% | 0.1545% | 3.878% | -1.34% | -1.073% | 1.382% | -4.65% | -7.373% | 0.7952% | -11.48% | 3.227% | -0.7531% | 2.61% | -1.767% | -35.94% | 0.8486% | 3.465% | 1.756% | 4.639%
| Z-Score (Absolute) |  | 1.04 | 4.13 | 1.01 | 0.09 | 0.92 | 0.49 | 0.42 | 0.25 | 1.38 | 2.12 | 0.09 | 3.23 | 0.74 | 0.33 | 0.58 | 0.61 | NA | 0.1 | 0.81 | 0.35 | 1.12
| Analytical results for Cu in OREAS 313 (Certified Value 0.036 wt.%) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0346 | 0.033934 | 0.035868 | 0.034886 | 0.03512 | 0.036556 | 0.0358 | 0.0352 | 0.035 | 0.036 | 0.0334 | 0.0376 | 0.0352 | 0.036 | 0.0357 | 0.03591 | 0.03669 | 0.03534 | 0.03632 | 0.0363 | 0.0356 | 0.0366 | 0.03594 | 0.03679 | 0.03319 | 1
| 1 | 2 | 0.0349 | 0.033169 | 0.0356 | 0.034654 | 0.03376 | 0.03685 | 0.0355 | 0.0357 | 0.0351 | 0.0365 | 0.0339 | 0.0372 | 0.0352 | 0.036704 | 0.0365 | 0.03641 | 0.037 | 0.03514 | 0.03624 | 0.0362 | 0.0355 | 0.036 | 0.0357 | 0.03699 | 0.0335 | 39
| 1 | 3 | 0.0346 | 0.034115 | 0.035502 | 0.034968 | 0.03408 | 0.035549 | 0.0355 | 0.0366 | 0.0353 | 0.0358 | 0.0342 | 0.0368 | 0.0363 | 0.036293 | 0.0455 | 0.03641 | 0.03723 | 0.03535 | 0.03647 | 0.0361 | 0.0354 | 0.0363 | 0.03623 | 0.03587 | 0.03306 | 16
| 1 | 4 | 0.035 | 0.032199 | 0.03549 | 0.035073 | 0.03578 | 0.036528 | 0.0354 | 0.0357 | 0.0359 | 0.0369 | 0.034 | 0.0365 | 0.035 | 0.03585 | 0.0361 | 0.03553 | 0.03628 | 0.03392 | 0.03635 | 0.0358 | 0.0356 | 0.0359 | 0.03555 | 0.03651 | 0.03382 | 0.035683
| 1 | 5 | 0.0349 | 0.03532 | 0.0358 | 0.035401 | 0.03419 | 0.035929 | 0.0359 | 0.0354 | 0.0357 | 0.0362 | 0.0345 | 0.0366 | 0.0355 | 0.036436 | 0.0357 | 0.03638 | 0.03722 | 0.03498 | 0.0345 | 0.0361 | 0.0354 | 0.0355 | 0.03566 | 0.03621 | 0.03338 | 18
| 1 | 6 | 0.0349 | 0.036219 | 0.0354 | 0.03494 | 0.03459 | 0.036549 | 0.0355 | 0.0358 | 0.0353 | 0.0357 | 0.034 | 0.0365 | 0.0352 | 0.036365 | 0.0398 | 0.03618 | 0.03766 | 0.03435 | 0.03654 | 0.0355 | 0.0356 | 0.0357 | 0.03605 | 0.03656 | 0.03319
| Mean |  | 0.034817 | 0.034159 | 0.03561 | 0.034987 | 0.034587 | 0.036327 | 0.0356 | 0.035733 | 0.035383 | 0.036183 | 0.034 | 0.036867 | 0.0354 | 0.036274 | 0.038217 | 0.036137 | 0.037013 | 0.034847 | 0.03607 | 0.036 | 0.035517 | 0.036 | 0.035855 | 0.036488 | 0.033357
| Median |  | 0.0349 | 0.034025 | 0.035551 | 0.034954 | 0.03439 | 0.036539 | 0.0355 | 0.0357 | 0.0353 | 0.0361 | 0.034 | 0.0367 | 0.0352 | 0.036329 | 0.0363 | 0.03628 | 0.03711 | 0.03506 | 0.036335 | 0.0361 | 0.03555 | 0.03595 | 0.03582 | 0.036535 | 0.033285
| Std Dev. |  | 0.00017224 | 0.0014468 | 0.00018597 | 0.00024595 | 0.00074912 | 0.00048559 | 0.0002 | 0.00048028 | 0.00034881 | 0.0004535 | 0.00036332 | 0.00044572 | 0.00046904 | 0.00030819 | 0.0038897 | 0.00035495 | 0.00047965 | 0.00058411 | 0.00077661 | 0.00029665 | 9.8319e-05 | 0.0004 | 0.00026113 | 0.00040202 | 0.00027544
| Rel.Std.Dev. |  | 0.4947% | 4.236% | 0.5222% | 0.703% | 2.166% | 1.337% | 0.5618% | 1.344% | 0.9858% | 1.253% | 1.069% | 1.209% | 1.325% | 0.8496% | 10.18% | 0.9822% | 1.296% | 1.676% | 2.153% | 0.824% | 0.2768% | 1.111% | 0.7283% | 1.102% | 0.8257%
| PDM3 |  | -2.427% | -4.269% | -0.2032% | -1.949% | -3.071% | 1.806% | -0.2312% | 0.1424% | -0.8384% | 1.404% | -4.715% | 3.319% | -0.7917% | 1.659% | 7.102% | 1.273% | 3.73% | -2.342% | 1.086% | 0.8898% | -0.4648% | 0.8898% | 0.4834% | 2.258% | -6.518%
| Z-Score (Absolute) |  | 1.14 | 1.96 | 0.15 | 0.93 | 1.43 | 0.74 | 0.17 | 0 | 0.44 | 0.56 | 2.16 | 1.41 | 0.42 | 0.67 | 3.09 | 0.5 | 1.6 | 1.11 | 0.42 | 0.33 | 0.27 | 0.33 | 0.15 | 0.94 | 2.96
| Analytical results for Dy in OREAS 313 (Certified Value 2.69 ppm) | [26 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
2
| 1 | 1 | 2.271 | 2.64 | 2.889 | 3.5 | 3.27 | 2.6 | 2.2 | [18 empty] | 1
| 1 | 2 | 2.328 | 2.6 | 3.1171 | 3.6 | 2.56 | 2.6 | 2.1 | [18 empty] | 5
| 1 | 3 | 2.3047 | 2.76 | 3.0121 | 3.5 | 2.63 | 2.7 | 2.1 | [18 empty] | 16
| 1 | 4 | 2.3006 | 2.6 | 3.0198 | 3.4 | 2.5 | 2.7 | 2.2 | [18 empty] | 2.6886
| 1 | 5 | 2.3881 | 2.62 | 2.9546 | 3.5 | 2.64 | 2.6 | 2.16 | [18 empty] | 19
| 1 | 6 | 2.3898 | 2.54 | 2.8839 | 3.6 | 2.68 | 2.7 | 2.09
| Mean |  | 2.3304 | 2.6267 | 2.9794 | 3.5167 | 2.7133 | 2.65 | 2.1417
| Median |  | 2.3163 | 2.61 | 2.9833 | 3.5 | 2.635 | 2.65 | 2.13
| Std Dev. |  | 0.048868 | 0.073394 | 0.088964 | 0.075277 | 0.28012 | 0.054772 | 0.051543
| Rel.Std.Dev. |  | 2.097% | 2.794% | 2.986% | 2.141% | 10.32% | 2.067% | 2.407%
| PDM3 |  | -13.32% | -2.303% | 10.82% | 30.8% | 0.9204% | -1.435% | -20.34%
| Z-Score (Absolute) |  | 0.67 | 0.05 | 0.69 | 1.83 | 0.13 | 0 | 1.07
| Analytical results for Er in OREAS 313 (Certified Value 1.34 ppm) | [26 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
2
| 1 | 1 | 1.2876 | 1.424 | 1.5242 | 2.2 | 1.33 | 1.4 | 1.08 | [18 empty] | 1
| 1 | 2 | 1.2416 | 1.412 | 1.5891 | 2.1 | 1.3 | 1.3 | 1.05 | [18 empty] | 6
| 1 | 3 | 1.2757 | 1.42 | 1.6385 | 2.1 | 1.4 | 1.3 | 1.03 | [18 empty] | 16
| 1 | 4 | 1.2473 | 1.42 | 1.6818 | 2.1 | 1.24 | 1.4 | 1.09 | [18 empty] | 1.3424
| 1 | 5 | 1.2592 | 1.436 | 1.499 | 2.2 | 1.52 | 1.4 | 1.07 | [18 empty] | 20
| 1 | 6 | 1.1968 | 1.43 | 1.5346 | 2.2 | 1.51 | 1.4 | 0.99
| Mean |  | 1.2514 | 1.4237 | 1.5779 | 2.15 | 1.3833 | 1.3667 | 1.0517
| Median |  | 1.2532 | 1.422 | 1.5618 | 2.15 | 1.365 | 1.4 | 1.06
| Std Dev. |  | 0.031799 | 0.0084301 | 0.07154 | 0.054772 | 0.11431 | 0.05164 | 0.037103
| Rel.Std.Dev. |  | 2.541% | 0.5921% | 4.534% | 2.548% | 8.263% | 3.779% | 3.528%
| PDM3 |  | -6.783% | 6.052% | 17.54% | 60.16% | 3.047% | 1.806% | -21.66%
| Z-Score (Absolute) |  | 0.67 | 0.21 | 0.99 | 3.92 | 0 | 0.09 | 1.69
| Analytical results for Eu in OREAS 313 (Certified Value 0.9 ppm) | [26 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [19 empty] | 2
3
| 1 | 1 | 0.9457 | 0.92175 | 0.96 | 0.91 | 0.9 | 0.83 | [19 empty] | 1
| 1 | 2 | 0.9541 | 0.95567 | 1.04 | 0.92 | 0.9 | 0.76 | [19 empty] | 7
| 1 | 3 | 0.967 | 0.94668 | 0.89 | 0.92 | 0.9 | 0.73 | [19 empty] | 16
| 1 | 4 | 0.9439 | 0.95809 | 0.93 | 0.91 | 0.9 | 0.79 | [19 empty] | 0.90384
| 1 | 5 | 0.921 | 0.91637 | 0.93 | 0.92 | 0.9 | 0.76 | [19 empty] | 21
| 1 | 6 | 0.914 | 0.91089 | 1.03 | 0.91 | 0.9 | 0.72
| Mean |  | 0.94095 | 0.93491 | 0.96333 | 0.915 | 0.9 | 0.765
| Median |  | 0.9448 | 0.93422 | 0.945 | 0.915 | 0.9 | 0.76
| Std Dev. |  | 0.020035 | 0.020981 | 0.059889 | 0.0054772 | 0 | 0.040373
| Rel.Std.Dev. |  | 2.129% | 2.244% | 6.217% | 0.5986% | 0% | 5.278%
| PDM3 |  | 4.106% | 3.438% | 6.582% | 1.235% | -0.4247% | -15.36%
| Z-Score (Absolute) |  | 0.2 | 0 | 0.96 | 0.67 | NA | 5.75
Results from laboratory 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 313 (Certified Value 6.15 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 6.07 | 6.335 | 6.26 | 6.03 | 6.06 | 6.2711 | 6.21 | 6.22 | 5.85 | 5.97 | 5.81 | 6.54 | 6.24 | 6.6803 | 5.83 | 6.2 | 5.9 | 5.9751 | 6.22 | 6.31 | 6.16 | 6.04 | 6.48 | 6.17 | 6.28 | 1
| 1 | 2 | 5.99 | 6.3183 | 6.28 | 6 | 5.77 | 6.0629 | 6.24 | 6.29 | 5.86 | 5.97 | 5.84 | 6.14 | 6.22 | 6.4976 | 6.16 | 6.23 | 6.02 | 6.0034 | 6.2 | 6.37 | 6.13 | 6.09 | 6.47 | 6.15 | 6.31 | #N/A
| 1 | 3 | 5.99 | 6.3223 | 6.2418 | 6.11 | 5.9 | 6.2011 | 6.23 | 6.45 | 5.86 | 6.03 | 5.94 | 6.15 | 6.44 | 6.5127 | 6.1 | 6.26 | 6.06 | 5.9152 | 6.17 | 6.32 | 6.12 | 6.08 | 6.55 | 6.06 | 6.25 | 16
| 1 | 4 | 6.06 | 6.3121 | 6.22 | 6.15 | 6.19 | 6.084 | 6.17 | 6.19 | 5.98 | 6.1 | 5.88 | 5.81 | 6.23 | 6.662 | 5.87 | 6.11 | 5.85 | 5.8477 | 6.22 | 6.12 | 6.15 | 6.04 | 6.37 | 6.44 | 6.41 | 6.1545
| 1 | 5 | 6 | 6.326 | 6.2212 | 6.09 | 5.88 | 6.109 | 6.34 | 6.07 | 5.92 | 5.96 | 5.95 | 6.12 | 6.3 | 6.569 | 6.23 | 6.37 | 5.98 | 5.9753 | 5.77 | 6.45 | 6.16 | 6.05 | 6.47 | 6.1 | 6.35 | 22
| 1 | 6 | 6.06 | 6.343 | 6.36 | 5.99 | 5.93 | 6.2607 | 6.24 | 6.34 | 5.87 | 5.9 | 5.94 | 6.22 | 6.23 | 6.5367 | 6.05 | 6.25 | 6.03 | 5.8627 | 6.29 | 6.29 | 6.16 | 6.16 | 6.49 | 6.06 | 6.18
| Mean |  | 6.0283 | 6.3261 | 6.2638 | 6.0617 | 5.955 | 6.1648 | 6.2383 | 6.26 | 5.89 | 5.9883 | 5.8933 | 6.1633 | 6.2767 | 6.5764 | 6.04 | 6.2367 | 5.9733 | 5.9299 | 6.145 | 6.31 | 6.1467 | 6.0767 | 6.4717 | 6.1633 | 6.2967
| Median |  | 6.03 | 6.3241 | 6.2509 | 6.06 | 5.915 | 6.155 | 6.235 | 6.255 | 5.865 | 5.97 | 5.91 | 6.145 | 6.235 | 6.5528 | 6.075 | 6.24 | 6 | 5.9451 | 6.21 | 6.315 | 6.155 | 6.065 | 6.475 | 6.125 | 6.295
| Std Dev. |  | 0.038687 | 0.011279 | 0.052438 | 0.064627 | 0.14816 | 0.091495 | 0.056362 | 0.13115 | 0.050596 | 0.068532 | 0.059217 | 0.23347 | 0.08501 | 0.077486 | 0.1595 | 0.084774 | 0.081894 | 0.0648 | 0.18791 | 0.10936 | 0.017512 | 0.045898 | 0.058109 | 0.14292 | 0.079917
| Rel.Std.Dev. |  | 0.6417% | 0.1783% | 0.8371% | 1.066% | 2.488% | 1.484% | 0.9035% | 2.095% | 0.859% | 1.144% | 1.005% | 3.788% | 1.354% | 1.178% | 2.641% | 1.359% | 1.371% | 1.093% | 3.058% | 1.733% | 0.2849% | 0.7553% | 0.8979% | 2.319% | 1.269%
| PDM3 |  | -2.05% | 2.788% | 1.776% | -1.509% | -3.242% | 0.167% | 1.362% | 1.714% | -4.298% | -2.7% | -4.244% | 0.1432% | 1.985% | 6.855% | -1.861% | 1.335% | -2.944% | -3.65% | -0.1547% | 2.526% | -0.1276% | -1.265% | 5.153% | 0.1432% | 2.31%
| Z-Score (Absolute) |  | 0.74 | 0.89 | 0.55 | 0.56 | 1.14 | 0.01 | 0.41 | 0.53 | 1.49 | 0.96 | 1.48 | 0 | 0.62 | 2.26 | 0.67 | 0.4 | 1.04 | 1.28 | 0.1 | 0.8 | 0.09 | 0.47 | 1.69 | 0 | 0.73
| Analytical results for Ga in OREAS 313 (Certified Value 14.3 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 13.2 | 15.1425 | 14.72 | 14.98 | 14.45 | 15.2746 | 13.4 | 13.7 | 13.45 | 13.3 | 13.7 | 14.6 | 12.06 | 16.5016 | 28.6 | 13.96 | 13.46 | 16 | 13.7 | 14.4 | 15.13 | 14.61 | 15.69 | 15 |  | 1
| 1 | 2 | 13.4 | 15.1073 | 14.46 | 14.97 | 13.55 | 15.5358 | 13.25 | 13.25 | 14.25 | 13.4 | 14 | 13.3 | 11.85 | 16.3284 | 35.6 | 14.45 | 13.46 | 16 | 13.4 | 15.2 | 14.1 | 14.78 | 15.84 | 16 |  | 21
| 1 | 3 | 13.5 | 15.105 | 15.04 | 14.96 | 14.16 | 14.6906 | 13.25 | 14 | 14.3 | 13.2 | 14.15 | 15.6 | 12.17 | 16.2657 | 32.2 | 14.45 | 13.45 | 15 | 13.2 | 14 | 15.76 | 14.55 | 15.82 | 16 |  | 16
| 1 | 4 | 13.5 | 15.0725 | 14.448 | 14.93 | 14.78 | 15.1547 | 12.9 | 13.6 | 14.05 | 13.35 | 13.75 | 14.4 | 12.06 | 16.4992 | 33.7 | 14.03 | 13.51 | 15 | 13.2 | 14.4 | 14.66 | 14.54 | 15.81 | 16 |  | 14.2576
| 1 | 5 | 13.5 | 15.0787 | 15.08 | 14.66 | 13.76 | 14.4838 | 13.45 | 13.45 | 14.7 | 12.9 | 14.2 | 15.6 | 12.17 | 16.4394 | 32.2 | 14.5 | 13.39 | 16 | 13.4 | 14.8 | 14.64 | 14.28 | 15.81 | 15 |  | 23
| 1 | 6 | 13.1 | 15.1387 | 14.26 | 14.94 | 14.13 | 14.8317 | 13.75 | 13.55 | 14.35 | 12.8 | 14.05 | 12.6 | 12.34 | 16.3807 | 34.4 | 14.5 | 13.58 | 15 | 13.1 | 14.6 | 15.5 | 14.61 | 15.6 | 16
| Mean |  | 13.3667 | 15.1075 | 14.668 | 14.9067 | 14.1383 | 14.9952 | 13.3333 | 13.5917 | 14.1833 | 13.1583 | 13.975 | 14.35 | 12.1083 | 16.4025 | 32.7833 | 14.315 | 13.475 | 15.5 | 13.3333 | 14.5667 | 14.965 | 14.5617 | 15.7617 | 15.6667
| Median |  | 13.45 | 15.1061 | 14.59 | 14.95 | 14.145 | 14.9932 | 13.325 | 13.575 | 14.275 | 13.25 | 14.025 | 14.5 | 12.115 | 16.41 | 32.95 | 14.45 | 13.46 | 15.5 | 13.3 | 14.5 | 14.895 | 14.58 | 15.81 | 16
| Std Dev. |  | 0.17512 | 0.029187 | 0.33733 | 0.12226 | 0.44692 | 0.39415 | 0.28048 | 0.25183 | 0.41673 | 0.24983 | 0.20676 | 1.2128 | 0.1629 | 0.095093 | 2.4318 | 0.24986 | 0.064109 | 0.54772 | 0.21602 | 0.40825 | 0.6152 | 0.16266 | 0.095376 | 0.5164
| Rel.Std.Dev. |  | 1.31% | 0.1932% | 2.3% | 0.8201% | 3.161% | 2.628% | 2.104% | 1.853% | 2.938% | 1.899% | 1.48% | 8.452% | 1.345% | 0.5797% | 7.418% | 1.745% | 0.4758% | 3.534% | 1.62% | 2.803% | 4.111% | 1.117% | 0.6051% | 3.296%
| PDM3 |  | -6.249% | 5.961% | 2.878% | 4.552% | -0.8366% | 5.173% | -6.483% | -4.671% | -0.521% | -7.71% | -1.982% | 0.648% | -15.07% | 15.04% | 129.9% | 0.4025% | -5.489% | 8.714% | -6.483% | 2.168% | 4.961% | 2.133% | 10.55% | 9.883%
| Z-Score (Absolute) |  | 0.93 | 0.74 | 0.32 | 0.55 | 0.19 | 0.64 | 0.96 | 0.71 | 0.14 | 1.13 | 0.34 | 0.02 | 2.14 | 1.99 | 17.73 | 0.02 | 0.82 | NA | 0.96 | 0.23 | 0.61 | 0.22 | 1.37 | NA
| Analytical results for Gd in OREAS 313 (Certified Value 3.86 ppm) | [26 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
2
| 1 | 1 | 4.7115 | 3.62 | 4.0596 | 4.4 | 3.71 | 3.9 | 2.86 | [18 empty] | 1
| 1 | 2 | 4.6758 | 3.58 | 4.3135 | 4.3 | 3.64 | 3.8 | 2.74 | [18 empty] | 8
| 1 | 3 | 4.675 | 3.58 | 4.1768 | 4.1 | 3.93 | 3.8 | 2.76 | [18 empty] | 16
| 1 | 4 | 4.698 | 3.42 | 4.2924 | 4.3 | 3.65 | 3.7 | 2.82 | [18 empty] | 3.8569
| 1 | 5 | 4.6042 | 3.6 | 3.9812 | 4.5 | 3.8 | 3.8 | 2.82 | [18 empty] | 24
| 1 | 6 | 4.6229 | 3.48 | 4.0179 | 4.2 | 3.79 | 3.8 | 2.76
| Mean |  | 4.6646 | 3.5467 | 4.1402 | 4.3 | 3.7533 | 3.8 | 2.7933
| Median |  | 4.6754 | 3.58 | 4.1182 | 4.3 | 3.75 | 3.8 | 2.79
| Std Dev. |  | 0.042275 | 0.078655 | 0.14232 | 0.14142 | 0.10967 | 0.063246 | 0.046762
| Rel.Std.Dev. |  | 0.9063% | 2.218% | 3.438% | 3.289% | 2.922% | 1.664% | 1.674%
| PDM3 |  | 20.94% | -8.043% | 7.347% | 11.49% | -2.685% | -1.475% | -27.58%
| Z-Score (Absolute) |  | 1.71 | 0.5 | 0.67 | 0.99 | 0.09 | 0 | 2
| Analytical results for Ge in OREAS 313 (Certified Value 0.14 ppm) | [26 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.23 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | 9.6 | 8.3 | < 0.2 | 0.12 | 0.14 | 0.14 | 0.14 | 0.19 | 0.3 | 0.1428 | 0.34 | 10.4 | 0.18 | 8 | < 0.05 | 10.7 | [9 empty] | 1
| 1 | 2 | 9.2 | 8.5 | < 0.2 | 0.11 | 0.12 | 0.14 | 0.09 | 0.18 | 0.2 | 0.155 | 0.35 | 10.7 | 0.08 | 8.6 | < 0.05 | 11.2 | [9 empty] | 23
| 1 | 3 | 9.2 | 8 | < 0.2 | 0.09 | 0.14 | 0.14 | 0.12 | 0.19 | 0.2 | 0.1246 | 0.35 | 10.6 | 0.09 | 8.2 | < 0.05 | 10.9 | [9 empty] | 16
| 1 | 4 | 9.4 | 8 | < 0.2 | 0.11 | 0.12 | 0.14 | 0.13 | 0.16 | < 0.1 | 0.1217 | 0.36 | 10.4 | 0.15 | 8.2 | < 0.05 | 11.3 | [9 empty] | 0.13559
| 1 | 5 | 9.3 | 8.3 | < 0.2 | 0.1 | 0.19 | 0.14 | 0.13 | 0.17 | 0.2 | 0.159 | 0.37 | 10.6 | 0.11 | 6.8 | < 0.05 | 10.8 | [9 empty] | 25
| 1 | 6 | 9.6 | 8.4 | < 0.2 | 0.11 | 0.14 | 0.14 | 0.13 | 0.17 | 0.2 | 0.1297 | 0.36 | 10.8 | 0.14 | 8.4 | < 0.05 | 10.9
| Mean |  | 9.3833 | 8.25 |  | 0.10667 | 0.14167 | 0.14 | 0.12333 | 0.17667 | 0.22 | 0.1388 | 0.355 | 10.5833 | 0.125 | 8.0333 |  | 10.9667
| Median |  | 9.35 | 8.3 |  | 0.11 | 0.14 | 0.14 | 0.13 | 0.175 | 0.2 | 0.13625 | 0.355 | 10.6 | 0.125 | 8.2 |  | 10.9
| Std Dev. |  | 0.18348 | 0.20736 |  | 0.010328 | 0.025626 | 0 | 0.017512 | 0.012111 | 0.044721 | 0.015896 | 0.010488 | 0.16021 | 0.038341 | 0.6377 |  | 0.23381
| Rel.Std.Dev. |  | 1.955% | 2.514% |  | 9.682% | 18.09% | 0% | 14.2% | 6.855% | 20.33% | 11.45% | 2.954% | 1.514% | 30.67% | 7.938% |  | 2.132%
| PDM3 |  | 6820% | 5984% |  | -21.33% | 4.481% | 3.252% | -9.04% | 30.29% | 62.25% | 2.367% | 161.8% | 7705% | -7.811% | 5825% |  | 7988%
| Z-Score (Absolute) |  | NA | NA | 0.81 | 0.67 | 0.06 | 0.03 | 0.32 | 0.79 | NA | 0 | 4.54 | 219.18 | 0.29 | NA | 2.39 | NA
Results from laboratories 2.01, 2.05, 2.12, 2.23 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 313 (Certified Value 2.08 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2 | 2 | 2.19 | 2.0678 | 2.3 | 2.1 | 2 | 2 | 1.7 | 2.2 | 2.04 | 2.2143 | 1.9 | 2.19 | 2.07 | 2.1 | 2.25 | 1.93 | 2.18 | 2 | 1.9 |  |  |  |  | 1
| 1 | 2 | 2.1 | 2.01 | 2.13 | 2.1197 | 2.4 | 2.2 | 2 | 2 | 1.8 | 2.4 | 2.12 | 2.3174 | 1.9 | 2.11 | 2.07 | 2 | 2.25 | 1.92 | 2.11 | 2 | 2 |  |  |  |  | 30
| 1 | 3 | 2.1 | 2.1 | 2.21 | 2.0889 | 2.3 | 2.1 | 1.9 | 2 | 1.8 | 2.3 | 2.17 | 2.21 | 1.9 | 2.07 | 2.04 | 2 | 2.18 | 1.92 | 2.12 | 2 | 2 |  |  |  |  | 16
| 1 | 4 | 2.1 | 2.01 | 2.18 | 2.1153 | 2.3 | 1.9 | 2 | 2 | 1.8 | 2.2 | 2.09 | 2.1817 | 1.8 | 2.14 | 1.99 | 2 | 2.23 | 1.85 | 2.16 | 2 | 2 |  |  |  |  | 2.0754
| 1 | 5 | 2 | 2.18 | 2.12 | 2.0361 | 2.3 | 2 | 1.9 | 2 | 1.8 | 2.3 | 2.13 | 2.2968 | 1.9 | 2.18 | 2.07 | 1.9 | 2.35 | 1.88 | 2.12 | 2 | 1.8 |  |  |  |  | 26
| 1 | 6 | 2 | 3.01 | 2.27 | 2.0112 | 2.4 | 2 | 2 | 2 | 1.8 | 2.3 | 2.2 | 2.2359 | 1.9 | 2.14 | 1.95 | 2.1 | 2.25 | 2.06 | 2.09 | 2 | 2
| Mean |  | 2.05 | 2.2183 | 2.1833 | 2.0732 | 2.3333 | 2.05 | 1.9667 | 2 | 1.7833 | 2.2833 | 2.125 | 2.2427 | 1.8833 | 2.1383 | 2.0317 | 2.0167 | 2.2517 | 1.9267 | 2.13 | 2 | 1.95
| Median |  | 2.05 | 2.055 | 2.185 | 2.0784 | 2.3 | 2.05 | 2 | 2 | 1.8 | 2.3 | 2.125 | 2.2251 | 1.9 | 2.14 | 2.055 | 2 | 2.25 | 1.92 | 2.12 | 2 | 2
| Std Dev. |  | 0.054772 | 0.39413 | 0.055015 | 0.043439 | 0.05164 | 0.10488 | 0.05164 | 0 | 0.040825 | 0.075277 | 0.056833 | 0.053198 | 0.040825 | 0.04446 | 0.050761 | 0.075277 | 0.055287 | 0.072019 | 0.033466 | 0 | 0.083666
| Rel.Std.Dev. |  | 2.672% | 17.77% | 2.52% | 2.095% | 2.213% | 5.116% | 2.626% | 0% | 2.289% | 3.297% | 2.674% | 2.372% | 2.168% | 2.079% | 2.498% | 3.733% | 2.455% | 3.738% | 1.571% | 0% | 4.291%
| PDM3 |  | -1.224% | 6.886% | 5.2% | -0.1084% | 12.43% | -1.224% | -5.24% | -3.634% | -14.07% | 10.02% | 2.389% | 8.06% | -9.255% | 3.032% | -2.108% | -2.831% | 8.492% | -7.167% | 2.63% | -3.634% | -6.043%
| Z-Score (Absolute) |  | 0.14 | 0.89 | 0.67 | 0 | 1.59 | 0.14 | 0.65 | NA | 1.77 | 1.29 | 0.32 | 1.04 | 1.16 | 0.4 | 0.25 | 0.35 | 1.09 | 0.9 | 0.35 | NA | 0.75
Results from laboratories 2.10 and 2.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 313 (Indicative Value < 1 ppm) | [26 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [23 empty] | 1
| No | No | 2.06 | 2.18 | [23 empty] | ppm
|  |  | 4A*MS | 4A*MS | [23 empty] | 2
2
| 1 | 1 | < 1 | < 0.01 | [23 empty] | 1
| 1 | 2 | < 1 | 0.18 | [23 empty] | 24
| 1 | 3 | < 1 | < 0.01 | [23 empty] | 16
| 1 | 4 | < 1 | < 0.01 | [23 empty] | < 1
| 1 | 5 | < 1 | < 0.01 | [23 empty] | 7
| 1 | 6 | < 1 | < 0.01
| Mean |  |  | 0.18
| Median |  |  | 0.18
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 313 (Certified Value 0.5 ppm) | [26 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
3
| 1 | 1 | 0.48672 | 0.48 | 0.53991 | 0.7 | 0.46 | 0.5 | 0.37 | [18 empty] | 1
| 1 | 2 | 0.4913 | 0.484 | 0.53538 | 0.7 | 0.46 | 0.5 | 0.39 | [18 empty] | 4
| 1 | 3 | 0.51584 | 0.5 | 0.54983 | 0.6 | 0.48 | 0.5 | 0.36 | [18 empty] | 16
| 1 | 4 | 0.46891 | 0.48 | 0.54405 | 0.6 | 0.48 | 0.5 | 0.39 | [18 empty] | 0.49743
| 1 | 5 | 0.48178 | 0.508 | 0.53968 | 0.7 | 0.48 | 0.5 | 0.37 | [18 empty] | 27
| 1 | 6 | 0.48412 | 0.49 | 0.52675 | 0.7 | 0.49 | 0.5 | 0.38
| Mean |  | 0.48811 | 0.49033 | 0.53927 | 0.66667 | 0.475 | 0.5 | 0.37667
| Median |  | 0.48542 | 0.487 | 0.5398 | 0.7 | 0.48 | 0.5 | 0.375
| Std Dev. |  | 0.015529 | 0.011483 | 0.0078333 | 0.05164 | 0.012247 | 0 | 0.012111
| Rel.Std.Dev. |  | 3.181% | 2.342% | 1.453% | 7.746% | 2.578% | 0% | 3.215%
| PDM3 |  | -1.874% | -1.427% | 8.41% | 34.02% | -4.51% | 0.516% | -24.28%
| Z-Score (Absolute) |  | 0.1 | 0 | 2.15 | 7.75 | 0.67 | 0.43 | 5
| Analytical results for In in OREAS 313 (Certified Value 0.29 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.27 | 0.28854 | 0.31 | < 1 | 0.31 | 0.29717 | 0.314 | 0.283 | 0.284 | 0.27 | 0.261 | 0.3 | 0.252 | 0.3342 | 0.305 | 0.3 | 0.27 | 0.3 | 0.23 | 0.25 | 0.28 | 0.31 | 0.296 | 0.29 |  | 1
| 1 | 2 | 0.27 | 0.2947 | 0.32 | < 1 | 0.3 | 0.30564 | 0.301 | 0.27 | 0.288 | 0.3 | 0.268 | 0.3 | 0.249 | 0.3237 | 0.308 | 0.28 | 0.3 | 0.3 | 0.21 | 0.26 | 0.28 | 0.29 | 0.297 | 0.29 |  | 11
| 1 | 3 | 0.27 | 0.3199 | 0.33 | < 1 | 0.3 | 0.30413 | 0.311 | 0.285 | 0.276 | 0.279 | 0.277 | 0.3 | 0.255 | 0.3371 | 0.306 | 0.32 | 0.28 | 0.3 | 0.2 | 0.25 | 0.29 | 0.29 | 0.291 | 0.3 |  | 16
| 1 | 4 | 0.27 | 0.29736 | 0.3 | < 1 | 0.31 | 0.3287 | 0.293 | 0.292 | 0.274 | 0.281 | 0.271 | 0.3 | 0.257 | 0.3007 | 0.287 | 0.29 | 0.3 | 0.3 | 0.21 | 0.26 | 0.3 | 0.3 | 0.3 | 0.3 |  | 0.29068
| 1 | 5 | 0.27 | 0.31584 | 0.32 | < 1 | 0.3 | 0.29748 | 0.311 | 0.263 | 0.287 | 0.264 | 0.258 | 0.3 | 0.258 | 0.3387 | 0.311 | 0.27 | 0.3 | 0.28 | 0.21 | 0.27 | 0.29 | 0.31 | 0.298 | 0.28 |  | 28
| 1 | 6 | 0.28 | 0.31514 | 0.3 | < 1 | 0.3 | 0.29411 | 0.314 | 0.281 | 0.284 | 0.258 | 0.277 | 0.3 | 0.258 | 0.2993 | 0.286 | 0.29 | 0.28 | 0.31 | 0.2 | 0.26 | 0.32 | 0.3 | 0.296 | 0.29
| Mean |  | 0.27167 | 0.30525 | 0.31333 |  | 0.30333 | 0.30454 | 0.30733 | 0.279 | 0.28217 | 0.27533 | 0.26867 | 0.3 | 0.25483 | 0.32228 | 0.3005 | 0.29167 | 0.28833 | 0.29833 | 0.21 | 0.25833 | 0.29333 | 0.3 | 0.29633 | 0.29167
| Median |  | 0.27 | 0.30625 | 0.315 |  | 0.3 | 0.3008 | 0.311 | 0.282 | 0.284 | 0.2745 | 0.2695 | 0.3 | 0.256 | 0.32895 | 0.3055 | 0.29 | 0.29 | 0.3 | 0.21 | 0.26 | 0.29 | 0.3 | 0.2965 | 0.29
| Std Dev. |  | 0.0040825 | 0.013247 | 0.012111 |  | 0.005164 | 0.012636 | 0.008501 | 0.010602 | 0.0058109 | 0.014909 | 0.0079666 | 0 | 0.003656 | 0.01804 | 0.011041 | 0.017224 | 0.013292 | 0.0098319 | 0.010954 | 0.0075277 | 0.015055 | 0.0089443 | 0.0030111 | 0.0075277
| Rel.Std.Dev. |  | 1.503% | 4.34% | 3.865% |  | 1.702% | 4.149% | 2.766% | 3.8% | 2.059% | 5.415% | 2.965% | 0% | 1.435% | 5.597% | 3.674% | 5.905% | 4.61% | 3.296% | 5.216% | 2.914% | 5.133% | 2.981% | 1.016% | 2.581%
| PDM3 |  | -6.541% | 5.011% | 7.793% |  | 4.353% | 4.768% | 5.729% | -4.018% | -2.929% | -5.279% | -7.573% | 3.206% | -12.33% | 10.87% | 3.378% | 0.3396% | -0.8071% | 2.633% | -27.76% | -11.13% | 0.913% | 3.206% | 1.945% | 0.3396%
| Z-Score (Absolute) |  | 1.23 | 0.67 | 1.13 | 11.7 | 0.57 | 0.63 | 0.79 | 0.81 | 0.63 | 1.02 | 1.4 | NA | 2.18 | 1.64 | 0.41 | 0.09 | 0.28 | 0.28 | 4.72 | 1.98 | 0 | 0.38 | 0.17 | 0.09
Results from laboratory 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 313 (Certified Value 2 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.88 | 2.0845 | 2 | 1.86 | 1.9922 | 2.0105 | 2.06 | 1.95 | 1.93 | 1.97 | 1.88 | 2.25 | 1.99 | 2.0707 | 1.94 | 1.9436 | 2.03 | 2.11 | 2.02 | 1.96 | 2 | 2.1017 | 2.04 | 2.01 |  | 1
| 1 | 2 | 1.9 | 2.047 | 2.04 | 1.84 | 1.8829 | 2.0195 | 2.05 | 1.98 | 1.95 | 1.98 | 1.89 | 2.28 | 2 | 2.0153 | 2.04 | 1.9644 | 2.06 | 2.11 | 2.04 | 1.95 | 2.01 | 2.1159 | 2.05 | 2.03 |  | #N/A
| 1 | 3 | 1.89 | 2.0742 | 2.0496 | 1.86 | 1.926 | 1.9971 | 2.07 | 2.03 | 1.93 | 1.97 | 1.92 | 2.03 | 2.04 | 2.0168 | 2 | 1.9715 | 2.07 | 2.12 | 2.03 | 1.94 | 2.01 | 2.1175 | 2.02 | 2.02 |  | 16
| 1 | 4 | 1.92 | 2.0419 | 2 | 1.86 | 2.0216 | 2.0464 | 2.06 | 1.97 | 2.02 | 2 | 1.91 | 1.95 | 1.98 | 2.0609 | 1.96 | 1.9215 | 1.97 | 2.11 | 1.98 | 1.95 | 2.01 | 2.0738 | 2.05 | 2.01 |  | 1.9987
| 1 | 5 | 1.89 | 2.0665 | 2.0139 | 1.87 | 1.9279 | 1.987 | 2.09 | 1.89 | 1.99 | 1.96 | 1.94 | 2 | 2 | 2.0519 | 2.06 | 1.9877 | 2.07 | 1.95 | 2.04 | 1.95 | 1.99 | 2.077 | 2.03 | 2.05 |  | 29
| 1 | 6 | 1.91 | 2.0667 | 2.02 | 1.85 | 1.9542 | 2.0462 | 2.07 | 2 | 1.95 | 1.95 | 1.93 | 2.35 | 1.99 | 2.0219 | 1.99 | 1.947 | 2.01 | 2.12 | 2.02 | 1.96 | 2.04 | 2.1105 | 2.02 | 2.03
| Mean |  | 1.8983 | 2.0635 | 2.0206 | 1.8567 | 1.9508 | 2.0178 | 2.0667 | 1.97 | 1.9617 | 1.9717 | 1.9117 | 2.1433 | 2 | 2.0396 | 1.9983 | 1.9559 | 2.035 | 2.0867 | 2.0217 | 1.9517 | 2.01 | 2.0994 | 2.035 | 2.025
| Median |  | 1.895 | 2.0666 | 2.017 | 1.86 | 1.9411 | 2.015 | 2.065 | 1.975 | 1.95 | 1.97 | 1.915 | 2.14 | 1.995 | 2.0369 | 1.995 | 1.9557 | 2.045 | 2.11 | 2.025 | 1.95 | 2.01 | 2.1061 | 2.035 | 2.025
| Std Dev. |  | 0.01472 | 0.016208 | 0.020549 | 0.010328 | 0.049976 | 0.024728 | 0.013663 | 0.047749 | 0.036009 | 0.017224 | 0.023166 | 0.16943 | 0.020976 | 0.024478 | 0.045789 | 0.023412 | 0.039875 | 0.067132 | 0.022286 | 0.0075277 | 0.016733 | 0.019419 | 0.013784 | 0.015166
| Rel.Std.Dev. |  | 0.7754% | 0.7855% | 1.017% | 0.5563% | 2.562% | 1.225% | 0.6611% | 2.424% | 1.836% | 0.8736% | 1.212% | 7.905% | 1.049% | 1.2% | 2.291% | 1.197% | 1.959% | 3.217% | 1.102% | 0.3857% | 0.8325% | 0.925% | 0.6773% | 0.7489%
| PDM3 |  | -5.021% | 3.242% | 1.096% | -7.105% | -2.395% | 0.9571% | 3.402% | -1.435% | -1.852% | -1.351% | -4.353% | 7.238% | 0.06618% | 2.047% | -0.01721% | -2.138% | 1.817% | 4.402% | 1.15% | -2.352% | 0.5665% | 5.039% | 1.817% | 1.317%
| Z-Score (Absolute) |  | 1.67 | 0.72 | 0.1 | 2.27 | 0.91 | 0.06 | 0.76 | 0.63 | 0.75 | 0.61 | 1.48 | 1.87 | 0.2 | 0.37 | 0.22 | 0.84 | 0.3 | 1.05 | 0.11 | 0.9 | 0.06 | 1.23 | 0.3 | 0.16
| Analytical results for La in OREAS 313 (Certified Value 28.9 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
1
| 1 | 1 | 29.18 | 28.2 | 25.31 | 31.22 | 32.5729 | 30.2 | 26.5 | 28 | 26.8 | 25.1 | 27.9 | 24.46 | 14.6 | 31.93 | 26.6 | 28.3 | 28 | 31.5 | 26.2 | 32.52 | 31.7 | 37 |  |  |  | 1
| 1 | 2 | 28.54 | 28.6 | 25.52 | 30.51 | 33.7028 | 31 | 26.9 | 27.9 | 26.5 | 25.5 | 25.5 | 23.74 | 15 | 32.23 | 26.1 | 27.8 | 29 | 31.2 | 25.7 | 31.31 | 30.7 | 37 |  |  |  | 12
| 1 | 3 | 28.74 | 29 | 26.53 | 31.83 | 32.4989 | 29.4 | 27 | 26.9 | 26.5 | 25.9 | 23.1 | 25.35 | 15 | 32.98 | 26 | 28.2 | 27 | 30.6 | 26.8 | 32.05 | 31.4 | 37 |  |  |  | 16
| 1 | 4 | 26.78 | 27.9 | 26.47 | 32.99 | 33.5381 | 30.8 | 28.6 | 28 | 27.5 | 25.3 | 22.3 | 24.78 | 14.5 | 31.92 | 25.8 | 27.5 | 27 | 31 | 26.3 | 31.46 | 31.3 | 37 |  |  |  | 28.8871
| 1 | 5 | 29.36 | 29 | 25.4 | 31.79 | 32.4551 | 31.2 | 25.6 | 28.6 | 25.7 | 26.1 | 25.4 | 26.49 | 15.1 | 32.85 | 25.9 | 26.1 | 28 | 32.9 | 25.7 | 31.56 | 31 | 38 |  |  |  | 30
| 1 | 6 | 28.1 | 27.4 | 25.74 | 32.88 | 32.7545 | 30.9 | 26.3 | 28.3 | 25.5 | 24.9 | 25.6 | 24.41 | 14.8 | 33.24 | 25.4 | 28.4 | 28 | 31.7 | 28.1 | 31.5 | 29.7 | 35
| Mean |  | 28.45 | 28.35 | 25.8283 | 31.87 | 32.9204 | 30.5833 | 26.8167 | 27.95 | 26.4167 | 25.4667 | 24.9667 | 24.8717 | 14.8333 | 32.525 | 25.9667 | 27.7167 | 27.8333 | 31.4833 | 26.4667 | 31.7333 | 30.9667 | 36.8333
| Median |  | 28.64 | 28.4 | 25.63 | 31.81 | 32.6637 | 30.85 | 26.7 | 28 | 26.5 | 25.4 | 25.45 | 24.62 | 14.9 | 32.54 | 25.95 | 28 | 28 | 31.35 | 26.25 | 31.53 | 31.15 | 37
| Std Dev. |  | 0.9343 | 0.63797 | 0.5402 | 0.95442 | 0.55427 | 0.67057 | 1.0068 | 0.57533 | 0.73326 | 0.46332 | 2.0037 | 0.95073 | 0.24221 | 0.57113 | 0.39328 | 0.8612 | 0.75277 | 0.79352 | 0.90037 | 0.4599 | 0.70899 | 0.98319
| Rel.Std.Dev. |  | 3.284% | 2.25% | 2.092% | 2.995% | 1.684% | 2.193% | 3.754% | 2.058% | 2.776% | 1.819% | 8.025% | 3.823% | 1.633% | 1.756% | 1.515% | 3.107% | 2.705% | 2.52% | 3.402% | 1.449% | 2.29% | 2.669%
| PDM3 |  | -1.513% | -1.859% | -10.59% | 10.33% | 13.96% | 5.872% | -7.167% | -3.244% | -8.552% | -11.84% | -13.57% | -13.9% | -48.65% | 12.59% | -10.11% | -4.052% | -3.648% | 8.988% | -8.379% | 9.853% | 7.199% | 27.51%
| Z-Score (Absolute) |  | 0.15 | 0.12 | 0.54 | 1.05 | 1.33 | 0.71 | 0.28 | 0.02 | 0.39 | 0.64 | 0.77 | 0.8 | 3.44 | 1.22 | 0.51 | 0.05 | 0.02 | 0.95 | 0.38 | 1.01 | 0.81 | 2.36
| Analytical results for Li in OREAS 313 (Certified Value 23.5 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 22.9 | 24.4745 | 20.2 | 24.17 | 22.6 | 23.3814 | 25 | 23.2 | 23.7 | 24.5 | 23.3 | 25 | 24.3 | 24.1422 | 20.7 | 22.6 | 23.5 | 24.8 | 19.6 | 22 | 21.7 | 24.4 | 24.3 | 23.5 |  | 1
| 1 | 2 | 22.6 | 24.4855 | 21.4 | 25.44 | 21.8 | 23.6397 | 25 | 23.8 | 23.7 | 24.2 | 23.7 | 24 | 24.5 | 23.1617 | 20.3 | 23.5 | 23.1 | 24.5 | 19.3 | 23 | 21.8 | 24.5 | 24.9 | 22.8 |  | #N/A
| 1 | 3 | 22.4 | 24.668 | 20.6 | 24.81 | 22.2 | 22.8273 | 24.9 | 24.4 | 23.7 | 24.4 | 23.8 | 24 | 24.9 | 22.1783 | 21.2 | 23.5 | 22.6 | 24.4 | 19.4 | 23 | 21.5 | 24.6 | 23.8 | 22.7 |  | 16
| 1 | 4 | 22.6 | 24.529 | 21.3 | 25.11 | 23.2 | 23.4871 | 24.8 | 23.5 | 24 | 24.8 | 23.4 | 24 | 24.6 | 22.1581 | 20.5 | 24.1 | 22.2 | 25 | 19.8 | 23 | 22.6 | 24.7 | 24.9 | 23.2 |  | 23.4667
| 1 | 5 | 22.6 | 24.6665 | 21 | 24.87 | 21.9 | 23.237 | 25.3 | 23.9 | 23.9 | 24.4 | 24.1 | 24 | 24.6 | 22.7887 | 22.7 | 23.4 | 22.4 | 22.7 | 19.1 | 22 | 22.4 | 24 | 23.7 | 21.7 |  | 31
| 1 | 6 | 22.4 | 24.535 | 21.5 | 24.81 | 22.5 | 24.2939 | 25.1 | 23.8 | 23.8 | 24.3 | 23.8 | 24 | 24.4 | 23.7378 | 22.1 | 23.7 | 22.7 | 25.3 | 18.7 | 22 | 23.4 | 24.4 | 24.7 | 21.8
| Mean |  | 22.5833 | 24.5597 | 21 | 24.8683 | 22.3667 | 23.4778 | 25.0167 | 23.7667 | 23.8 | 24.4333 | 23.6833 | 24.1667 | 24.55 | 23.0278 | 21.25 | 23.4667 | 22.75 | 24.45 | 19.3167 | 22.5 | 22.2333 | 24.4333 | 24.3833 | 22.6167
| Median |  | 22.6 | 24.532 | 21.15 | 24.84 | 22.35 | 23.4343 | 25 | 23.8 | 23.75 | 24.4 | 23.75 | 24 | 24.55 | 22.9752 | 20.95 | 23.5 | 22.65 | 24.65 | 19.35 | 22.5 | 22.1 | 24.45 | 24.5 | 22.75
| Std Dev. |  | 0.18348 | 0.086548 | 0.5099 | 0.41945 | 0.5164 | 0.48638 | 0.17224 | 0.40332 | 0.12649 | 0.20656 | 0.29269 | 0.40825 | 0.20736 | 0.81229 | 0.95864 | 0.49261 | 0.47645 | 0.91815 | 0.38687 | 0.54772 | 0.71181 | 0.24221 | 0.53821 | 0.73052
| Rel.Std.Dev. |  | 0.8125% | 0.3524% | 2.428% | 1.687% | 2.309% | 2.072% | 0.6885% | 1.697% | 0.5315% | 0.8454% | 1.236% | 1.689% | 0.8447% | 3.527% | 4.511% | 2.099% | 2.094% | 3.755% | 2.003% | 2.434% | 3.202% | 0.9913% | 2.207% | 3.23%
| PDM3 |  | -3.764% | 4.658% | -10.51% | 5.973% | -4.688% | 0.0472% | 6.605% | 1.278% | 1.42% | 4.119% | 0.9232% | 2.983% | 4.616% | -1.87% | -9.446% | -5.694e-05% | -3.054% | 4.19% | -17.68% | -4.119% | -5.256% | 4.119% | 3.906% | -3.622%
| Z-Score (Absolute) |  | 0.73 | 0.72 | 1.9 | 0.95 | 0.89 | 0.08 | 1.06 | 0.14 | 0.16 | 0.63 | 0.08 | 0.43 | 0.71 | 0.41 | 1.71 | 0.08 | 0.61 | 0.64 | 3.14 | 0.79 | 0.99 | 0.63 | 0.59 | 0.71
| Analytical results for Lu in OREAS 313 (Certified Value 0.21 ppm) | [26 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.13 | 2.16 | 2.17 | 2.20 | 2.21 | [15 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [15 empty] | 2
2
| 1 | 1 | 0.1598 | 0.212 | < 1 | 0.26953 | 0.3 | 0.18 | 0.21 | 0.2 | 0.18 | 0.28 | [15 empty] | 1
| 1 | 2 | 0.1496 | 0.216 | < 1 | 0.26408 | 0.3 | 0.18 | 0.21 | 0.2 | 0.17 | 0.28 | [15 empty] | 5
| 1 | 3 | 0.1528 | 0.226 | < 1 | 0.25747 | 0.3 | 0.19 | 0.22 | 0.2 | 0.18 | 0.26 | [15 empty] | 16
| 1 | 4 | 0.1422 | 0.211 | < 1 | 0.27233 | 0.3 | 0.18 | 0.21 | 0.2 | 0.17 | 0.26 | [15 empty] | 0.21098
| 1 | 5 | 0.1422 | 0.224 | < 1 | 0.25502 | 0.2 | 0.18 | 0.21 | 0.2 | 0.18 | 0.29 | [15 empty] | 32
| 1 | 6 | 0.1683 | 0.21 | < 1 | 0.24883 | 0.3 | 0.2 | 0.22 | 0.2 | 0.17 | 0.27
| Mean |  | 0.15248 | 0.2165 |  | 0.26121 | 0.28333 | 0.185 | 0.21333 | 0.2 | 0.175 | 0.27333
| Median |  | 0.1512 | 0.214 |  | 0.26078 | 0.3 | 0.18 | 0.21 | 0.2 | 0.175 | 0.275
| Std Dev. |  | 0.010233 | 0.006921 |  | 0.0090193 | 0.040825 | 0.0083666 | 0.005164 | 3.0405e-17 | 0.0054772 | 0.012111
| Rel.Std.Dev. |  | 6.711% | 3.197% |  | 3.453% | 14.41% | 4.522% | 2.421% | 1.52e-14% | 3.13% | 4.431%
| PDM3 |  | -27.73% | 2.616% |  | 23.81% | 34.29% | -12.31% | 1.116% | -5.204% | -17.05% | 29.55%
| Z-Score (Absolute) |  | 0.98 | 0.02 | 4.46 | 0.72 | NA | 0.47 | 0.02 | NA | 0.62 | 0.91
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 313 (Certified Value 1.03 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 0.984 | 1.0627 | 1.024 | 1.03 | 1.0151 | 1.1128 | 1.05 | 0.98 | 0.98 | 0.99 | 0.96 | 1.2 | 1.02 | 1.0663 | 0.99 | 0.9651 | 1.04 | 1.008 | 1.04 | 0.97 | 1.02 | 1 | 1.0564 | 1.02 | 1.05 | 1
| 1 | 2 | 0.995 | 1.0726 | 1.03 | 1.03 | 0.9574 | 1.1194 | 1.05 | 1 | 0.98 | 0.98 | 0.97 | 1.13 | 1.019 | 1.036 | 1.07 | 0.9649 | 1.06 | 1.0066 | 1.04 | 1.06 | 1.01 | 1.01 | 1.054 | 1.02 | 1.06 | #N/A
| 1 | 3 | 0.988 | 1.0983 | 1.0353 | 1.05 | 0.9809 | 1.0909 | 1.05 | 1.03 | 0.99 | 0.99 | 0.99 | 1.12 | 1.048 | 1.0382 | 1.06 | 0.9763 | 1.06 | 0.9922 | 1.04 | 0.98 | 1.02 | 1.01 | 1.0624 | 1 | 1.05 | 16
| 1 | 4 | 1 | 1.0585 | 1.022 | 1.02 | 1.0323 | 1.1163 | 1.04 | 0.99 | 0.99 | 1.02 | 0.97 | 1.03 | 1.014 | 1.0701 | 0.99 | 0.9559 | 1.04 | 0.9842 | 1.05 | 0.98 | 1.01 | 1 | 1.041 | 1.06 | 1.06 | 1.0263
| 1 | 5 | 0.99 | 1.0967 | 1.0311 | 1.03 | 0.9698 | 1.0979 | 1.07 | 0.95 | 0.99 | 0.99 | 0.98 | 1.06 | 1.031 | 1.049 | 1.07 | 0.9728 | 1.06 | 1.0094 | 0.97 | 1.06 | 1.01 | 1 | 1.0522 | 1.01 | 1.06 | 33
| 1 | 6 | 0.998 | 1.0791 | 1.038 | 1.05 | 0.9799 | 1.1306 | 1.05 | 1.01 | 0.98 | 0.98 | 0.98 | 1.12 | 1.019 | 1.0367 | 1.02 | 0.9759 | 1.06 | 0.988 | 1.05 | 1.06 | 1.03 | 1.02 | 1.0497 | 1 | 1.04
| Mean |  | 0.9925 | 1.078 | 1.0301 | 1.035 | 0.98923 | 1.1113 | 1.0517 | 0.99333 | 0.985 | 0.99167 | 0.975 | 1.11 | 1.0252 | 1.0494 | 1.0333 | 0.96848 | 1.0533 | 0.99807 | 1.0317 | 1.0183 | 1.0167 | 1.0067 | 1.0526 | 1.0183 | 1.0533
| Median |  | 0.9925 | 1.0759 | 1.0306 | 1.03 | 0.9804 | 1.1146 | 1.05 | 0.995 | 0.985 | 0.99 | 0.975 | 1.12 | 1.0195 | 1.0436 | 1.04 | 0.96895 | 1.06 | 0.9994 | 1.04 | 1.02 | 1.015 | 1.005 | 1.0531 | 1.015 | 1.055
| Std Dev. |  | 0.0061887 | 0.016798 | 0.0062173 | 0.012247 | 0.028536 | 0.014564 | 0.0098319 | 0.027325 | 0.0054772 | 0.01472 | 0.010488 | 0.05933 | 0.012513 | 0.015366 | 0.038297 | 0.007967 | 0.010328 | 0.011207 | 0.030605 | 0.045789 | 0.008165 | 0.008165 | 0.0071494 | 0.022286 | 0.008165
| Rel.Std.Dev. |  | 0.6235% | 1.558% | 0.6036% | 1.183% | 2.885% | 1.311% | 0.9349% | 2.751% | 0.5561% | 1.484% | 1.076% | 5.345% | 1.221% | 1.464% | 3.706% | 0.8226% | 0.9805% | 1.123% | 2.967% | 4.497% | 0.8031% | 0.8111% | 0.6792% | 2.188% | 0.7752%
| PDM3 |  | -3.291% | 5.038% | 0.3698% | 0.8505% | -3.609% | 8.286% | 2.474% | -3.21% | -4.022% | -3.372% | -4.996% | 8.158% | -0.1077% | 2.252% | 0.6881% | -5.631% | 2.637% | -2.748% | 0.5257% | -0.7735% | -0.9359% | -1.91% | 2.567% | -0.7735% | 2.637%
| Z-Score (Absolute) |  | 0.8 | 1.3 | 0.12 | 0.24 | 0.88 | 2.12 | 0.65 | 0.78 | 0.99 | 0.82 | 1.23 | 2.08 | 0 | 0.59 | 0.2 | 1.39 | 0.69 | 0.67 | 0.16 | 0.17 | 0.21 | 0.45 | 0.67 | 0.17 | 0.69
| Analytical results for Mn in OREAS 313 (Certified Value 0.186 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.175 | 0.19395 | 0.19 | 0.1877 | 0.1844 | 0.18997 | 0.189 | 0.1815 | 0.1795 | 0.182 | 0.171 | 0.197 | 0.1883 | 0.2049 | 0.179 | 0.1944 | 0.1773 | 0.1925 | 0.1917 | 0.1773 | 0.1742 | 0.1783 | 0.1926 | 0.1903 | 0.1941 | 1
| 1 | 2 | 0.177 | 0.19221 | 0.18563 | 0.1856 | 0.1778 | 0.19247 | 0.1905 | 0.185 | 0.178 | 0.188 | 0.172 | 0.184 | 0.187 | 0.19694 | 0.188 | 0.1941 | 0.1798 | 0.1926 | 0.1873 | 0.1791 | 0.1745 | 0.1763 | 0.1907 | 0.1882 | 0.1949 | 35
| 1 | 3 | 0.175 | 0.1935 | 0.1919 | 0.1887 | 0.1784 | 0.18484 | 0.19 | 0.1895 | 0.18 | 0.185 | 0.175 | 0.182 | 0.1937 | 0.2051 | 0.185 | 0.1949 | 0.1803 | 0.1897 | 0.1884 | 0.1777 | 0.1746 | 0.1803 | 0.1938 | 0.1855 | 0.1947 | 16
| 1 | 4 | 0.178 | 0.19373 | 0.18707 | 0.1895 | 0.1858 | 0.19147 | 0.1865 | 0.18 | 0.1765 | 0.1885 | 0.172 | 0.178 | 0.1881 | 0.19232 | 0.18 | 0.1902 | 0.1759 | 0.1884 | 0.1911 | 0.1764 | 0.1749 | 0.1777 | 0.1892 | 0.1907 | 0.1993 | 0.18567
| 1 | 5 | 0.176 | 0.19312 | 0.18715 | 0.189 | 0.1806 | 0.18368 | 0.1945 | 0.1775 | 0.1795 | 0.185 | 0.175 | 0.178 | 0.1897 | 0.19282 | 0.19 | 0.1973 | 0.1811 | 0.1926 | 0.1764 | 0.1792 | 0.1748 | 0.1767 | 0.1893 | 0.1868 | 0.1979 | 34
| 1 | 6 | 0.177 | 0.1914 | 0.18354 | 0.1864 | 0.1813 | 0.18979 | 0.189 | 0.188 | 0.1795 | 0.182 | 0.1745 | 0.189 | 0.1879 | 0.2043 | 0.185 | 0.1937 | 0.1817 | 0.189 | 0.192 | 0.1771 | 0.1742 | 0.1754 | 0.1909 | 0.1872 | 0.1902
| Mean |  | 0.17633 | 0.19298 | 0.18755 | 0.18782 | 0.18138 | 0.1887 | 0.18992 | 0.18358 | 0.17883 | 0.18508 | 0.17325 | 0.18467 | 0.18912 | 0.1994 | 0.1845 | 0.1941 | 0.17935 | 0.1908 | 0.18782 | 0.1778 | 0.17453 | 0.17745 | 0.19108 | 0.18812 | 0.19518
| Median |  | 0.1765 | 0.19331 | 0.18711 | 0.1882 | 0.18095 | 0.18988 | 0.1895 | 0.18325 | 0.1795 | 0.185 | 0.17325 | 0.183 | 0.1882 | 0.20062 | 0.185 | 0.19425 | 0.18005 | 0.1911 | 0.18975 | 0.1775 | 0.17455 | 0.1772 | 0.1908 | 0.1877 | 0.1948
| Std Dev. |  | 0.0012111 | 0.00098761 | 0.0030028 | 0.0015459 | 0.0031928 | 0.0036012 | 0.0026347 | 0.0047055 | 0.0013292 | 0.0028003 | 0.0017819 | 0.0073121 | 0.0024087 | 0.0061028 | 0.0043243 | 0.0022952 | 0.0022713 | 0.0019789 | 0.0059044 | 0.0011278 | 0.00029439 | 0.0017318 | 0.0018214 | 0.002043 | 0.0031827
| Rel.Std.Dev. |  | 0.6868% | 0.5118% | 1.601% | 0.8231% | 1.76% | 1.908% | 1.387% | 2.563% | 0.7432% | 1.513% | 1.028% | 3.96% | 1.274% | 3.061% | 2.344% | 1.182% | 1.266% | 1.037% | 3.144% | 0.6343% | 0.1687% | 0.9759% | 0.9532% | 1.086% | 1.631%
| PDM3 |  | -5.026% | 3.942% | 1.014% | 1.159% | -2.306% | 1.637% | 2.29% | -1.121% | -3.68% | -0.3134% | -6.687% | -0.5378% | 1.859% | 7.395% | -0.6276% | 4.543% | -3.401% | 2.766% | 1.159% | -4.236% | -5.996% | -4.425% | 2.918% | 1.32% | 5.126%
| Z-Score (Absolute) |  | 1.91 | 0.92 | 0 | 0.05 | 1.05 | 0.2 | 0.4 | 0.67 | 1.48 | 0.42 | 2.43 | 0.49 | 0.27 | 2.02 | 0.52 | 1.11 | 1.39 | 0.55 | 0.05 | 1.66 | 2.21 | 1.72 | 0.6 | 0.1 | 1.3
| Analytical results for Mo in OREAS 313 (Certified Value 5.26 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 5.2 | < 5 | 5.8 | 5.34 | 5.5 | 5.0897 | 5.5 | 5.13 | 4.94 | 5 | 5.01 | 6 | 4.74 | 5.0737 | 5.41 | 5.3 | 5.43 | 6 | 5.5 | 5.1 | 5.1 | 4.66 | 5.2 | 5.81 | 5.2 | 1
| 1 | 2 | 5.1 | < 5 | 5.7 | 5.2 | 5.3 | 5.2263 | 5.8 | 4.96 | 4.91 | 5.07 | 5.88 | 5 | 4.78 | 5.1719 | 5.57 | 5.5 | 5.66 | 6.2 | 5.2 | 4.7 | 5.13 | 4.64 | 5.3 | 6.03 | 5.4 | 14
| 1 | 3 | 5.1 | 5.21 | 5.6 | 5.28 | 5.4 | 5.0094 | 5.5 | 5.12 | 4.9 | 4.98 | 5.18 | 6 | 4.97 | 4.91 | 5.58 | 5.4 | 5.07 | 5.5 | 5.2 | 5 | 5.09 | 5.3 | 5.4 | 5.8 | 5.4 | 16
| 1 | 4 | 5.1 | 5.39 | 5.4 | 5.51 | 5.4 | 5.2006 | 5.44 | 5.08 | 4.81 | 5.13 | 5.05 | 6 | 4.73 | 5.2325 | 5.25 | 5.3 | 5.39 | 5.7 | 5.3 | 4.8 | 5.11 | 4.82 | 5.2 | 5.96 | 5.3 | 5.2571
| 1 | 5 | 5.1 | 5.32 | 5.6 | 5.08 | 5.3 | 4.8435 | 5.45 | 5.05 | 5.05 | 4.82 | 5.34 | 5 | 4.79 | 5.333 | 5.45 | 5.6 | 5.41 | 5.9 | 4.9 | 4.7 | 5.03 | 4.84 | 5.3 | 6 | 5.3 | 35
| 1 | 6 | 5.2 | 5.51 | 5.6 | 5.16 | 5.4 | 5.1241 | 5.6 | 4.91 | 5.01 | 4.72 | 5.15 | 7 | 4.95 | 4.9528 | 5.5 | 5.6 | 5.39 | 5.6 | 5.4 | 5.2 | 5.19 | 5.3 | 5.3 | 5.78 | 5.4
| Mean |  | 5.1333 | 5.3575 | 5.6167 | 5.2617 | 5.3833 | 5.0822 | 5.5483 | 5.0417 | 4.9367 | 4.9533 | 5.2683 | 5.8333 | 4.8267 | 5.1123 | 5.46 | 5.45 | 5.3917 | 5.8167 | 5.25 | 4.9167 | 5.1083 | 4.9267 | 5.2833 | 5.8967 | 5.3333
| Median |  | 5.1 | 5.355 | 5.6 | 5.24 | 5.4 | 5.1069 | 5.5 | 5.065 | 4.925 | 4.99 | 5.165 | 6 | 4.785 | 5.1228 | 5.475 | 5.45 | 5.4 | 5.8 | 5.25 | 4.9 | 5.105 | 4.83 | 5.3 | 5.885 | 5.35
| Std Dev. |  | 0.05164 | 0.1258 | 0.13292 | 0.15184 | 0.075277 | 0.14062 | 0.13571 | 0.088863 | 0.085245 | 0.15488 | 0.32109 | 0.75277 | 0.10596 | 0.16405 | 0.12231 | 0.13784 | 0.1883 | 0.26394 | 0.20736 | 0.2137 | 0.052313 | 0.30031 | 0.075277 | 0.11219 | 0.08165
| Rel.Std.Dev. |  | 1.006% | 2.348% | 2.366% | 2.886% | 1.398% | 2.767% | 2.446% | 1.763% | 1.727% | 3.127% | 6.095% | 12.9% | 2.195% | 3.209% | 2.24% | 2.529% | 3.492% | 4.538% | 3.95% | 4.346% | 1.024% | 6.096% | 1.425% | 1.903% | 1.531%
| PDM3 |  | -2.354% | 1.91% | 6.84% | 0.08687% | 2.401% | -3.326% | 5.54% | -4.098% | -6.095% | -5.778% | 0.2137% | 10.96% | -8.188% | -2.754% | 3.86% | 3.669% | 2.56% | 10.64% | -0.1351% | -6.476% | -2.83% | -6.285% | 0.499% | 12.17% | 1.45%
| Z-Score (Absolute) |  | 0.41 | 2.88 | 1.22 | 0.02 | 0.43 | 0.59 | 0.99 | 0.73 | 1.08 | 1.02 | 0.04 | NA | 1.45 | 0.49 | 0.69 | 0.66 | 0.46 | 1.9 | 0.02 | 1.15 | 0.5 | 1.11 | 0.09 | 2.17 | 0.26
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 313 (Certified Value 0.076 wt.%) | [26 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.073 | 0.14539 | 0.1479 | 0.1 | 0.073 | 0.07824 | 0.08 | 0.07 | 0.07 | 0.12 | 0.08 | 0.08 | 0.077 | 0.1147 | 0.07 | 0.0714 | 0.083 | 0.0614 | 0.077 | 0.08 | 0.08 | 0.075 | 0.0725 | 0.08 | 0.08 | 1
| 1 | 2 | 0.071 | 0.14856 | 0.1428 | 0.08 | 0.0656 | 0.079008 | 0.08 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.077 | 0.1164 | 0.07 | 0.0709 | 0.081 | 0.0615 | 0.076 | 0.08 | 0.08 | 0.076 | 0.0732 | 0.08 | 0.09 | 6
| 1 | 3 | 0.073 | 0.14565 | 0.1424 | 0.1 | 0.068 | 0.076202 | 0.08 | 0.08 | 0.07 | 0.11 | 0.08 | 0.08 | 0.08 | 0.1203 | 0.08 | 0.072 | 0.082 | 0.0602 | 0.076 | 0.08 | 0.08 | 0.075 | 0.0734 | 0.08 | 0.08 | 16
| 1 | 4 | 0.073 | 0.14631 | 0.1428 | 0.13 | 0.0747 | 0.080441 | 0.09 | 0.07 | 0.07 | 0.12 | 0.08 | 0.07 | 0.077 | 0.1176 | 0.07 | 0.0708 | 0.078 | 0.0602 | 0.078 | 0.08 | 0.08 | 0.075 | 0.0718 | 0.09 | 0.08 | 0.075703
| 1 | 5 | 0.074 | 0.1468 | 0.1482 | 0.12 | 0.0663 | 0.075649 | 0.08 | 0.07 | 0.07 | 0.13 | 0.08 | 0.07 | 0.077 | 0.1219 | 0.07 | 0.0721 | 0.084 | 0.0612 | 0.071 | 0.08 | 0.08 | 0.074 | 0.0723 | 0.08 | 0.09 | 36
| 1 | 6 | 0.072 | 0.14666 | 0.1484 | 0.09 | 0.0681 | 0.077542 | 0.08 | 0.07 | 0.07 | 0.11 | 0.08 | 0.08 | 0.078 | 0.1133 | 0.07 | 0.0724 | 0.083 | 0.06 | 0.078 | 0.08 | 0.08 | 0.075 | 0.0721 | 0.08 | 0.09
| Mean |  | 0.072667 | 0.14656 | 0.14542 | 0.10333 | 0.069283 | 0.077847 | 0.081667 | 0.071667 | 0.07 | 0.11167 | 0.08 | 0.076667 | 0.077667 | 0.11737 | 0.071667 | 0.0716 | 0.081833 | 0.06075 | 0.076 | 0.08 | 0.08 | 0.075 | 0.07255 | 0.081667 | 0.085
| Median |  | 0.073 | 0.14648 | 0.14535 | 0.1 | 0.06805 | 0.077891 | 0.08 | 0.07 | 0.07 | 0.115 | 0.08 | 0.08 | 0.077 | 0.117 | 0.07 | 0.0717 | 0.0825 | 0.0607 | 0.0765 | 0.08 | 0.08 | 0.075 | 0.0724 | 0.08 | 0.085
| Std Dev. |  | 0.0010328 | 0.0011249 | 0.0030202 | 0.018619 | 0.0037059 | 0.0017811 | 0.0040825 | 0.0040825 | 0 | 0.017224 | 0 | 0.005164 | 0.0012111 | 0.00328 | 0.0040825 | 0.00066633 | 0.002137 | 0.00068629 | 0.0026077 | 0 | 0 | 0.00063246 | 0.00062849 | 0.0040825 | 0.0054772
| Rel.Std.Dev. |  | 1.421% | 0.7675% | 2.077% | 18.02% | 5.349% | 2.288% | 4.999% | 5.696% | 0% | 15.42% | 0% | 6.736% | 1.559% | 2.795% | 5.696% | 0.9306% | 2.611% | 1.13% | 3.431% | 0% | 0% | 0.8433% | 0.8663% | 4.999% | 6.444%
| PDM3 |  | -4.011% | 93.6% | 92.09% | 36.5% | -8.48% | 2.832% | 7.877% | -5.332% | -7.534% | 47.51% | 5.676% | 1.273% | 2.594% | 55.04% | -5.332% | -5.42% | 8.098% | -19.75% | 0.3921% | 5.676% | 5.676% | -0.9289% | -4.165% | 7.877% | 12.28%
| Z-Score (Absolute) |  | 0.66 | 8.75 | 8.6 | 3.24 | 1.09 | 0 | 0.49 | 0.79 | 1 | 4.31 | 0.27 | 0.15 | 0.02 | 5.03 | 0.79 | 0.8 | 0.51 | 2.18 | 0.24 | 0.27 | 0.27 | 0.36 | 0.67 | 0.49 | 0.91
| Analytical results for Nb in OREAS 313 (Certified Value 3.3 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 3 | 3.1292 | 3.3 | 6.92 | 3.77 | 3.8509 | 4.1 | 3.9 | 3 | 3.2 | 3.3 | 8.6 | 2.4 | 3.4288 | 3.3 | 4.32 | 3.19 | 3.3 | < 5 | 2.7 | 3.43 | 2.98 | 5.4 | 2.9 |  | 1
| 1 | 2 | 2.9 | 3.12 | 3.4 | 7.77 | 3.18 | 4.0097 | 4 | 3.8 | 3.2 | 3.4 | 3.6 | 8.5 | 2.5 | 3.4738 | 3.3 | 3.56 | 2.98 | 3.3 | < 5 | 2.6 | 3.58 | 2.88 | 5.6 | 2.6 |  | 16
| 1 | 3 | 2.9 | 3.1027 | 3 | 7.83 | 3.67 | 3.9061 | 3.8 | 3.9 | 3.2 | 3.6 | 3.4 | 7.9 | 2.6 | 3.7629 | 3.4 | 3.29 | 2.82 | 3.2 | < 5 | 2.8 | 3.85 | 3.02 | 5.5 | 2.9 |  | 16
| 1 | 4 | 2.7 | 3.1551 | 3.2 | 7.47 | 3.73 | 3.9154 | 4.1 | 3.9 | 2.9 | 3.8 | 3.5 | 7.9 | 2.5 | 3.7709 | 2.6 | 4.2 | 3.07 | 3.3 | < 5 | 2.8 | 3.31 | 2.92 | 5.6 | 2.7 |  | 3.3011
| 1 | 5 | 2.9 | 3.1647 | 3.1 | 5.8 | 3.22 | 3.7792 | 3.9 | 3.7 | 2.2 | 3.3 | 3.3 | 8.4 | 2.8 | 3.3784 | 3.3 | 4.03 | 3.02 | 3 | < 5 | 2.5 | 3.35 | 2.95 | 5.6 | 2.6 |  | 37
| 1 | 6 | 2.9 | 3.1035 | 3.1 | 5.4 | 3.57 | 3.7871 | 4 | 4 | 3.4 | 3.2 | 3.5 | 8.4 | 2.6 | 3.2594 | 3.3 | 4.08 | 3.27 | 3.3 | < 5 | 2.7 | 3.8 | 2.96 | 5.5 | 2.8
| Mean |  | 2.8833 | 3.1292 | 3.1833 | 6.865 | 3.5233 | 3.8747 | 3.9833 | 3.8667 | 2.9833 | 3.4167 | 3.4333 | 8.2833 | 2.5667 | 3.5124 | 3.2 | 3.9133 | 3.0583 | 3.2333 |  | 2.6833 | 3.5533 | 2.9517 | 5.5333 | 2.75
| Median |  | 2.9 | 3.1246 | 3.15 | 7.195 | 3.62 | 3.8785 | 4 | 3.9 | 3.1 | 3.35 | 3.45 | 8.4 | 2.55 | 3.4513 | 3.3 | 4.055 | 3.045 | 3.3 |  | 2.7 | 3.505 | 2.955 | 5.55 | 2.75
| Std Dev. |  | 0.098319 | 0.025994 | 0.1472 | 1.0392 | 0.25967 | 0.087442 | 0.1169 | 0.10328 | 0.42151 | 0.24014 | 0.12111 | 0.30605 | 0.13663 | 0.20977 | 0.29665 | 0.40058 | 0.15917 | 0.12111 |  | 0.1169 | 0.23036 | 0.048339 | 0.08165 | 0.13784
| Rel.Std.Dev. |  | 3.41% | 0.8307% | 4.624% | 15.14% | 7.37% | 2.257% | 2.935% | 2.671% | 14.13% | 7.028% | 3.527% | 3.695% | 5.323% | 5.972% | 9.27% | 10.24% | 5.205% | 3.746% |  | 4.357% | 6.483% | 1.638% | 1.476% | 5.012%
| PDM3 |  | -12.66% | -5.209% | -3.569% | 108% | 6.731% | 17.38% | 20.67% | 17.13% | -9.627% | 3.499% | 4.004% | 150.9% | -22.25% | 6.398% | -3.064% | 18.54% | -7.355% | -2.054% |  | -18.72% | 7.639% | -10.59% | 67.62% | -16.7%
| Z-Score (Absolute) |  | 0.73 | 0.32 | 0.23 | 5.86 | 0.33 | 0.91 | 1.09 | 0.9 | 0.57 | 0.15 | 0.18 | 8.2 | 1.25 | 0.31 | 0.21 | 0.97 | 0.44 | 0.15 | 1.37 | 1.06 | 0.38 | 0.62 | 3.65 | 0.95
| Analytical results for Nd in OREAS 313 (Certified Value 26.6 ppm) | [26 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.18 | 2.20 | [17 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [17 empty] | 1
2
| 1 | 1 | 23.79 | 25.8 | 28.0249 | 27.5 | 26.96 | 27.2 | 12 | 21.6 | [17 empty] | 1
| 1 | 2 | 23.718 | 26.2 | 29.7573 | 26.9 | 27.06 | 27.6 | 12 | 20.1 | [17 empty] | 7
| 1 | 3 | 23.693 | 26.4 | 27.8744 | 26.1 | 27.72 | 27.4 | 11 | 20.7 | [17 empty] | 16
| 1 | 4 | 23.648 | 25.8 | 28.9545 | 26.1 | 26.58 | 26.1 | 11 | 21.2 | [17 empty] | 26.5575
| 1 | 5 | 23.714 | 26.2 | 27.4048 | 28.6 | 27.74 | 27.1 | 12 | 21 | [17 empty] | 38
| 1 | 6 | 23.816 | 25.6 | 27.6652 | 27.3 | 27.35 | 26.7 | 11 | 20.6
| Mean |  | 23.7298 | 26 | 28.2802 | 27.0833 | 27.235 | 27.0167 | 11.5 | 20.8667
| Median |  | 23.716 | 26 | 27.9497 | 27.1 | 27.205 | 27.15 | 11.5 | 20.85
| Std Dev. |  | 0.062432 | 0.30984 | 0.89566 | 0.94745 | 0.4558 | 0.54191 | 0.54772 | 0.52026
| Rel.Std.Dev. |  | 0.2631% | 1.192% | 3.167% | 3.498% | 1.674% | 2.006% | 4.763% | 2.493%
| PDM3 |  | -10.65% | -2.099% | 6.487% | 1.98% | 2.551% | 1.729% | -56.7% | -21.43%
| Z-Score (Absolute) |  | 2.18 | 0.67 | 0.84 | 0.04 | 0.14 | 0 | NA | 4.08
| Analytical results for Ni in OREAS 313 (Certified Value 15.9 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 15 | 14.99 | 16.4 | 16.26 | 15.7 | 15.1514 | 15.9 | 14.2 | 15 | 15 | 15.5 | 17 | 16 | 17.5857 | 16.6 | 15.9 | 15.8 | 17.1 | 17.1 | 12.8 | 15.3 | 15.5 | 17 | 17.6 | 14.9 | 1
| 1 | 2 | 15 | 15.63 | 16.6 | 15.91 | 14.6 | 15.5273 | 19.1 | 13.2 | 15.6 | 14.8 | 17.2 | 17 | 16 | 17.3355 | 22.7 | 16.7 | 16.4 | 17 | 16.5 | 12.9 | 15.9 | 16.5 | 16.9 | 17.8 | 15.7 | 18
| 1 | 3 | 16 | 15.77 | 17.9 | 15.47 | 15.3 | 15.0345 | 15.3 | 14.2 | 16.1 | 14.8 | 16.4 | 19 | 17 | 17.4499 | 19.2 | 16.3 | 15.5 | 15.5 | 15.8 | 12.1 | 15.7 | 15.9 | 16 | 17.9 | 14.9 | 16
| 1 | 4 | 16 | 14.44 | 16.5 | 15.6 | 15.9 | 15.3789 | 15.4 | 13.4 | 15.5 | 16.4 | 16 | 18 | 16 | 17.7188 | 18.8 | 16.5 | 15.6 | 16.1 | 16.4 | 13.1 | 15.7 | 15.8 | 16.5 | 18 | 15.6 | 15.9264
| 1 | 5 | 16 | 15.84 | 17.4 | 16.04 | 14.8 | 14.4601 | 15.7 | 14 | 15.8 | 14.5 | 17 | 16 | 15 | 17.3677 | 18.7 | 16.4 | 16 | 17 | 15 | 12.6 | 15.7 | 16 | 16.3 | 18 | 15.3 | 39
| 1 | 6 | 16 | 15.26 | 16.4 | 15.32 | 16.3 | 15.749 | 16.4 | 14 | 16.4 | 14.5 | 17.4 | 17 | 16 | 17.7551 | 18.9 | 16 | 16.2 | 15.8 | 16.2 | 12.1 | 15.3 | 16.6 | 16 | 17.7 | 15.1
| Mean |  | 15.6667 | 15.3217 | 16.8667 | 15.7667 | 15.4333 | 15.2169 | 16.3 | 13.8333 | 15.7333 | 15 | 16.5833 | 17.3333 | 16 | 17.5355 | 19.15 | 16.3 | 15.9167 | 16.4167 | 16.1667 | 12.6 | 15.6 | 16.05 | 16.45 | 17.8333 | 15.25
| Median |  | 16 | 15.445 | 16.55 | 15.755 | 15.5 | 15.2652 | 15.8 | 14 | 15.7 | 14.8 | 16.7 | 17 | 16 | 17.5178 | 18.85 | 16.35 | 15.9 | 16.55 | 16.3 | 12.7 | 15.7 | 15.95 | 16.4 | 17.85 | 15.2
| Std Dev. |  | 0.5164 | 0.53961 | 0.6314 | 0.36164 | 0.65625 | 0.45078 | 1.4269 | 0.4274 | 0.48854 | 0.71274 | 0.74409 | 1.0328 | 0.63246 | 0.17879 | 1.9746 | 0.30332 | 0.34881 | 0.70261 | 0.71181 | 0.41952 | 0.24495 | 0.42308 | 0.43243 | 0.1633 | 0.34496
| Rel.Std.Dev. |  | 3.296% | 3.522% | 3.743% | 2.294% | 4.252% | 2.962% | 8.754% | 3.09% | 3.105% | 4.752% | 4.487% | 5.958% | 3.953% | 1.02% | 10.31% | 1.861% | 2.191% | 4.28% | 4.403% | 3.33% | 1.57% | 2.636% | 2.629% | 0.9157% | 2.262%
| PDM3 |  | -1.631% | -3.797% | 5.904% | -1.003% | -3.096% | -4.455% | 2.346% | -13.14% | -1.212% | -5.817% | 4.125% | 8.834% | 0.4623% | 10.1% | 20.24% | 2.346% | -0.06091% | 3.079% | 1.509% | -20.89% | -2.049% | 0.7763% | 3.288% | 11.97% | -4.247%
| Z-Score (Absolute) |  | NA | 0.78 | 1.04 | 0.25 | 0.64 | 0.9 | 0.37 | 2.52 | 0.29 | 1.15 | 0.7 | NA | NA | 1.82 | 3.71 | 0.37 | 0.08 | 0.51 | 0.21 | 3.97 | 0.45 | 0.08 | 0.55 | 2.17 | 0.86
Results from laboratories 2.01, 2.12 and 2.13 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 313 (Certified Value 0.031 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.031 | 0.0318 | 0.03 | 0.0315 | 0.030903 | 0.033 | 0.03 | 0.03 | 0.032 | 0.03 | 0.033 | 0.032 | 0.0314 | 0.034 | 0.0289 | 0.031 | 0.0293 | 0.031 | 0.0321 | 0.03 | 0.0306 | 0.0324 | 0.0313 | 0.031 |  | 1
| 1 | 2 | 0.031 | 0.032 | 0.03 | 0.0282 | 0.032615 | 0.032 | 0.031 | 0.03 | 0.032 | 0.03 | 0.033 | 0.032 | 0.0319 | 0.035 | 0.0302 | 0.031 | 0.0301 | 0.031 | 0.0315 | 0.03 | 0.0308 | 0.0342 | 0.0312 | 0.032 |  | #N/A
| 1 | 3 | 0.031 | 0.0323 | 0.03 | 0.0284 | 0.03146 | 0.032 | 0.032 | 0.03 | 0.032 | 0.031 | 0.033 | 0.033 | 0.0319 | 0.034 | 0.0295 | 0.032 | 0.0295 | 0.031 | 0.031 | 0.03 | 0.0308 | 0.0346 | 0.0307 | 0.031 |  | 16
| 1 | 4 | 0.032 | 0.0318 | 0.03 | 0.0315 | 0.03191 | 0.032 | 0.03 | 0.032 | 0.032 | 0.03 | 0.033 | 0.033 | 0.0311 | 0.033 | 0.0295 | 0.031 | 0.0294 | 0.032 | 0.0313 | 0.03 | 0.0306 | 0.0315 | 0.0307 | 0.032 |  | 0.031266
| 1 | 5 | 0.031 | 0.03192 | 0.03 | 0.0292 | 0.030311 | 0.033 | 0.031 | 0.031 | 0.032 | 0.031 | 0.032 | 0.032 | 0.0324 | 0.033 | 0.0304 | 0.032 | 0.0287 | 0.029 | 0.0315 | 0.03 | 0.0305 | 0.0328 | 0.0302 | 0.032 |  | 40
| 1 | 6 | 0.031 | 0.0316 | 0.03 | 0.03 | 0.032146 | 0.033 | 0.031 | 0.03 | 0.032 | 0.031 | 0.032 | 0.033 | 0.0318 | 0.033 | 0.0295 | 0.031 | 0.0288 | 0.032 | 0.0317 | 0.03 | 0.031 | 0.0336 | 0.0309 | 0.031
| Mean |  | 0.031167 | 0.031903 | 0.03 | 0.0298 | 0.031557 | 0.0325 | 0.030833 | 0.0305 | 0.032 | 0.0305 | 0.032667 | 0.0325 | 0.03175 | 0.033667 | 0.029667 | 0.031333 | 0.0293 | 0.031 | 0.031517 | 0.03 | 0.030717 | 0.033183 | 0.030833 | 0.0315
| Median |  | 0.031 | 0.03186 | 0.03 | 0.0296 | 0.031685 | 0.0325 | 0.031 | 0.03 | 0.032 | 0.0305 | 0.033 | 0.0325 | 0.03185 | 0.0335 | 0.0295 | 0.031 | 0.02935 | 0.031 | 0.0315 | 0.03 | 0.0307 | 0.0332 | 0.0308 | 0.0315
| Std Dev. |  | 0.00040825 | 0.00023678 | 0 | 0.0014629 | 0.00084613 | 0.00054772 | 0.00075277 | 0.00083666 | 0 | 0.00054772 | 0.0005164 | 0.00054772 | 0.00045056 | 0.0008165 | 0.0005465 | 0.0005164 | 0.0005099 | 0.0010954 | 0.00037103 | 0 | 0.00018348 | 0.0011669 | 0.00039833 | 0.00054772
| Rel.Std.Dev. |  | 1.31% | 0.7422% | 0% | 4.909% | 2.681% | 1.685% | 2.441% | 2.743% | 0% | 1.796% | 1.581% | 1.685% | 1.419% | 2.425% | 1.842% | 1.648% | 1.74% | 3.534% | 1.177% | 0% | 0.5973% | 3.517% | 1.292% | 1.739%
| PDM3 |  | -0.319% | 2.037% | -4.05% | -4.69% | 0.9306% | 3.945% | -1.385% | -2.451% | 2.346% | -2.451% | 4.478% | 3.945% | 1.547% | 7.677% | -5.117% | 0.214% | -6.289% | -0.8521% | 0.8004% | -4.05% | -1.758% | 6.131% | -1.385% | 0.7471%
| Z-Score (Absolute) |  | 0.07 | 0.59 | 1.12 | 1.3 | 0.28 | 1.12 | 0.37 | 0.67 | 0.67 | 0.67 | 1.27 | 1.12 | 0.45 | 2.17 | 1.42 | 0.07 | 1.75 | 0.22 | 0.24 | 1.12 | 0.48 | 1.74 | 0.37 | 0.22
| Analytical results for Pb in OREAS 313 (Certified Value 3.9 wt.%) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*MS | 4A*OES | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 3.75 | 4.0181 | 4 | 3.46 | 3.818 | 3.8285 | 3.89 | 3.85 | 3.89 | 3.88 | 3.88 | 3.92 | 3.91 | 4.0432 | 3.7 | 3.9162 | - | 3.7696 | 3.9 | 3.92 | 4.09 | 4.05 | 3.9761 | 3.78 | 4.01 | 1
| 1 | 2 | 3.57 | 3.9772 | 3.92 | 3.51 | 3.8851 | 3.7996 | 3.88 | 3.88 | 3.86 | 3.89 | 3.86 | 3.99 | 3.84 | 4.0456 | 3.6 | 3.8339 | - | 3.7853 | 3.85 | 3.91 | 3.96 | 4.11 | 3.9307 | 3.79 | 3.99 | 40
| 1 | 3 | 3.65 | 4.0032 | 3.9 | 3.57 | 3.8706 | 3.8405 | 3.89 | 3.89 | 3.88 | 3.91 | 3.88 | 4.11 | 3.85 | 3.9069 | 3.71 | 3.9097 | - | 3.7843 | 3.9 | 3.88 | 3.92 | 4.09 | 3.9501 | 3.76 | 3.99 | 16
| 1 | 4 | 3.56 | 3.9986 | 3.9964 | 3.49 | 3.8468 | 3.8369 | 3.89 | 3.84 | 3.99 | 3.88 | 3.87 | 4.04 | 3.8 | 3.9568 | 3.46 | 3.9106 | - | 3.7477 | 3.91 | 3.95 | 4.09 | 4.05 | 4.0203 | 3.77 | 3.97 | 3.8963
| 1 | 5 | 3.68 | 3.9714 | 3.96 | 3.48 | 3.8754 | 3.8474 | 3.93 | 3.88 | 3.86 | 3.87 | 3.83 | 4.05 | 3.63 | 3.9584 | 3.98 | 4.0004 | - | 3.8179 | 3.97 | 3.89 | 4.04 | 3.99 | 3.9616 | 3.82 | 4.03 | 41
| 1 | 6 | 3.71 | 3.9461 | 3.977 | 3.46 | 3.8945 | 3.8605 | 3.93 | 3.85 | 3.88 | 3.89 | 3.88 | 4.07 | 3.81 | 3.9503 | 3.71 | 3.9287 | - | 3.7059 | 3.87 | 3.92 | 4.06 | 4.07 | 4.0545 | 3.79 | 4
| Mean |  | 3.6533 | 3.9858 | 3.9589 | 3.495 | 3.8651 | 3.8356 | 3.9017 | 3.865 | 3.8933 | 3.8867 | 3.8667 | 4.03 | 3.8067 | 3.9769 | 3.6933 | 3.9166 |  | 3.7685 | 3.9 | 3.9117 | 4.0267 | 4.06 | 3.9822 | 3.785 | 3.9983
| Median |  | 3.665 | 3.9879 | 3.9685 | 3.485 | 3.873 | 3.8387 | 3.89 | 3.865 | 3.88 | 3.885 | 3.875 | 4.045 | 3.825 | 3.9576 | 3.705 | 3.9134 |  | 3.777 | 3.9 | 3.915 | 4.05 | 4.06 | 3.9688 | 3.785 | 3.995
| Std Dev. |  | 0.07607 | 0.025935 | 0.041004 | 0.041352 | 0.028121 | 0.020631 | 0.022286 | 0.020736 | 0.048854 | 0.013663 | 0.019664 | 0.066633 | 0.094798 | 0.055626 | 0.17084 | 0.053088 |  | 0.038275 | 0.040988 | 0.024833 | 0.070899 | 0.041473 | 0.046526 | 0.020736 | 0.020412
| Rel.Std.Dev. |  | 2.082% | 0.6507% | 1.036% | 1.183% | 0.7276% | 0.5379% | 0.5712% | 0.5365% | 1.255% | 0.3515% | 0.5085% | 1.653% | 2.49% | 1.399% | 4.626% | 1.355% |  | 1.016% | 1.051% | 0.6348% | 1.761% | 1.021% | 1.168% | 0.5479% | 0.5105%
| PDM3 |  | -6.235% | 2.297% | 1.607% | -10.3% | -0.8011% | -1.558% | 0.1383% | -0.8028% | -0.07563% | -0.2467% | -0.76% | 3.432% | -2.3% | 2.068% | -5.209% | 0.5211% |  | -3.281% | 0.09548% | 0.3949% | 3.346% | 4.202% | 2.206% | -2.856% | 2.619%
| Z-Score (Absolute) |  | 1.98 | 0.73 | 0.51 | 3.27 | 0.26 | 0.5 | 0.04 | 0.26 | 0.03 | 0.08 | 0.24 | 1.08 | 0.73 | 0.65 | 1.65 | 0.16 | NA | 1.04 | 0.03 | 0.12 | 1.06 | 1.33 | 0.7 | 0.91 | 0.83
| Analytical results for Pr in OREAS 313 (Certified Value 6.83 ppm) | [26 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
2
| 1 | 1 | 5.8638 | 6.82 | 7.9318 | 7.8 | 7.17 | 7 | 5.94 | [18 empty] | 1
| 1 | 2 | 5.9117 | 6.96 | 8.2745 | 7.1 | 7.13 | 7 | 5.79 | [18 empty] | 19
| 1 | 3 | 5.8758 | 6.688 | 7.8452 | 7.3 | 7.39 | 6.8 | 5.66 | [18 empty] | 16
| 1 | 4 | 5.9049 | 6.71 | 8.055 | 7.2 | 7 | 6.5 | 6.08 | [18 empty] | 6.831
| 1 | 5 | 5.8782 | 6.726 | 7.7888 | 7.5 | 7.42 | 6.9 | 5.99 | [18 empty] | 42
| 1 | 6 | 5.8008 | 6.66 | 7.9181 | 6.9 | 7.26 | 6.7 | 6
| Mean |  | 5.8725 | 6.7607 | 7.9689 | 7.3 | 7.2283 | 6.8167 | 5.91
| Median |  | 5.877 | 6.718 | 7.925 | 7.25 | 7.215 | 6.85 | 5.965
| Std Dev. |  | 0.039614 | 0.11175 | 0.17466 | 0.31623 | 0.16068 | 0.19408 | 0.15569
| Rel.Std.Dev. |  | 0.6746% | 1.653% | 2.192% | 4.332% | 2.223% | 2.847% | 2.634%
| PDM3 |  | -14.03% | -1.03% | 16.66% | 6.865% | 5.816% | -0.2104% | -13.48%
| Z-Score (Absolute) |  | 1.32 | 0.08 | 1.61 | 0.67 | 0.57 | 0 | 1.26
| Analytical results for Rb in OREAS 313 (Certified Value 108 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 0
0
| 1 | 1 | 100 | 105.09 | 105.9 | 106.38 | 113.1896 | 111.5 | 104 | 106 | 109.5 | 97.6 | 128 | 95.3 | 117.1765 | 105 | 112.06 | 105.1 | 112.2 | 103 | 107 | 110.68 | 118.7 | 114 |  |  |  | 1
| 1 | 2 | 101 | 105.336 | 107.7 | 103.64 | 116.6192 | 110 | 104.5 | 103.5 | 111 | 100.5 | 112 | 96.1 | 115.0036 | 108 | 114.27 | 105.8 | 109.7 | 107 | 109 | 110.53 | 121.7 | 115.7 |  |  |  | 20
| 1 | 3 | 99.3 | 105.728 | 106.8 | 104.76 | 111.2493 | 111 | 110.5 | 102.5 | 111.5 | 102 | 116 | 98.8 | 115.4969 | 108 | 115.76 | 105.6 | 110.7 | 103 | 109 | 111.3 | 119.5 | 116.2 |  |  |  | 16
| 1 | 4 | 99.8 | 104.548 | 103.3 | 109.62 | 113.3669 | 110 | 105.5 | 101.5 | 114 | 99.6 | 114 | 97.6 | 118.1656 | 104 | 109.69 | 100.6 | 114.3 | 103 | 107 | 110.94 | 122.2 | 115.8 |  |  |  | 107.6525
| 1 | 5 | 100 | 105.6 | 107 | 103.89 | 110.6474 | 110 | 106 | 107 | 108 | 103.5 | 117 | 97.7 | 117.9235 | 108 | 114.72 | 105.3 | 106.6 | 108 | 105 | 112.08 | 122.8 | 111.1 |  |  |  | 43
| 1 | 6 | 101 | 104.638 | 104.5 | 106.87 | 113.1112 | 112.5 | 103.5 | 106 | 106.5 | 101 | 109 | 98.1 | 115.4677 | 107 | 113.69 | 100.2 | 113.8 | 105 | 109 | 110.18 | 121.1 | 111.2
| Mean |  | 100.1833 | 105.1567 | 105.8667 | 105.86 | 113.0306 | 110.8333 | 105.6667 | 104.4167 | 110.0833 | 100.7 | 116 | 97.2667 | 116.539 | 106.6667 | 113.365 | 103.7667 | 111.2167 | 104.8333 | 107.6667 | 110.9517 | 121 | 114
| Median |  | 100 | 105.213 | 106.35 | 105.57 | 113.1504 | 110.5 | 105 | 104.75 | 110.25 | 100.75 | 115 | 97.65 | 116.3367 | 107.5 | 113.98 | 105.2 | 111.45 | 104 | 108 | 110.81 | 121.4 | 114.85
| Std Dev. |  | 0.6824 | 0.48971 | 1.6717 | 2.256 | 2.0917 | 1.0328 | 2.543 | 2.223 | 2.6724 | 2.0239 | 6.5422 | 1.3095 | 1.3828 | 1.7512 | 2.1774 | 2.622 | 2.8646 | 2.2286 | 1.633 | 0.66958 | 1.595 | 2.3332
| Rel.Std.Dev. |  | 0.6811% | 0.4657% | 1.579% | 2.131% | 1.851% | 0.9318% | 2.407% | 2.129% | 2.428% | 2.01% | 5.64% | 1.346% | 1.187% | 1.642% | 1.921% | 2.527% | 2.576% | 2.126% | 1.517% | 0.6035% | 1.318% | 2.047%
| PDM3 |  | -6.938% | -2.318% | -1.659% | -1.665% | 4.996% | 2.955% | -1.845% | -3.006% | 2.258% | -6.458% | 7.754% | -9.648% | 8.255% | -0.9158% | 5.306% | -3.61% | 3.311% | -2.619% | 0.01314% | 3.065% | 12.4% | 5.896%
| Z-Score (Absolute) |  | 1.26 | 0.36 | 0.24 | 0.24 | 1.06 | 0.66 | 0.27 | 0.5 | 0.53 | 1.17 | 1.6 | 1.79 | 1.7 | 0.09 | 1.12 | 0.62 | 0.73 | 0.42 | 0.09 | 0.69 | 2.5 | 1.24
| Analytical results for Re in OREAS 313 (Certified Value < 0.002 ppm) | [26 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.005 | < 0.01 | 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.001 | < 0.05 | 0.002 | 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [8 empty] | 1
| 1 | 2 | < 0.05 | < 0.005 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.001 | < 0.05 | < 0.002 | 0.003 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [8 empty] | 27
| 1 | 3 | < 0.05 | < 0.005 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | 0.002 | 0.003 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [8 empty] | 16
| 1 | 4 | < 0.05 | < 0.005 | < 0.01 | < 0.002 | 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.001 | < 0.05 | < 0.002 |  | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [8 empty] | < 0.002
| 1 | 5 | < 0.05 | < 0.005 | < 0.01 | < 0.002 | 0.003 | < 0.002 | < 0.002 | 0.003 | < 0.001 | < 0.05 | < 0.002 | 0.003 | < 0.002 | < 0.003 | < 0.005 | 0.002 | < 0.005 | [8 empty] | 44
| 1 | 6 | < 0.05 | < 0.005 | < 0.01 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005
| Mean |  |  |  |  | 0.002 | 0.0023333 |  | 0.002 | 0.0023333 | 0.0015 |  | 0.002 | 0.0024 | 0.003 |  |  | 0.002
| Median |  |  |  |  | 0.002 | 0.002 |  | 0.002 | 0.002 | 0.0015 |  | 0.002 | 0.002 | 0.003 |  |  | 0.002
| Std Dev. | [5 empty] | 0.00057735 |  |  | 0.00057735 | 0.00070711 |  |  | 0.00054772
| Rel.Std.Dev. | [5 empty] | 24.74% |  |  | 24.74% | 47.14% |  |  | 22.82%
PDM3
| Z-Score (Absolute) |  | 23.6 | 0.84 | 3.37 | 0.51 | 0 | 0.67 | 0.51 | 0 | 0.84 | 23.6 | 0.51 | 0.74 | 0.34 | 0.17 | 0.84 | 0.51 | 0.84
| Analytical results for S in OREAS 313 (Certified Value 4.44 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 4.42 | 4.9897 | 4.48 | 6.59 | 4.58 | 4.6074 | 4.62 | 4.37 | 4.3 | 4.41 | 4.08 | 4.62 | 4.329 | 4.6242 | 4.79 | 4.11 | 4.52 | 4.3033 | 4.4 | 4.13 | 4.26 | 4.46 | 4.69 | 4.48 | 4.11 | 1
| 1 | 2 | 4.5 | 4.9863 | 4.34 | 6.98 | 4.5 | 4.6288 | 4.62 | 4.41 | 4.35 | 4.39 | 4.12 | 4.62 | 4.316 | 4.6315 | 4.86 | 4.22 | 4.56 | 4.1839 | 4.3 | 4.05 | 4.26 | 4.48 | 4.72 | 4.4 | 4.07 | #N/A
| 1 | 3 | 4.46 | 4.9661 | 4.4 | 6.62 | 4.53 | 4.7783 | 4.6 | 4.55 | 4.37 | 4.46 | 4.18 | 4.59 | 4.487 | 4.5183 | 4.85 | 4.2 | 4.57 | 4.1467 | 4.3 | 4.17 | 4.29 | 4.52 | 4.61 | 4.41 | 4.07 | 16
| 1 | 4 | 4.52 | 4.9698 | 4.4064 | 6.61 | 4.43 | 4.6339 | 4.57 | 4.36 | 4.36 | 4.51 | 4.07 | 4.63 | 4.395 | 4.5391 | 4.65 | 4.2 | 4.49 | 4.1717 | 4.3 | 4.01 | 4.26 | 4.53 | 4.63 | 4.53 | 4.14 | 4.4423
| 1 | 5 | 4.49 | 4.9589 | 4.36 | 6.77 | 4.47 | 4.557 | 4.69 | 4.4 | 4.43 | 4.41 | 4.2 | 4.55 | 4.329 | 4.4989 | 4.69 | 4.23 | 4.59 | 4.1898 | 4 | 4.16 | 4.26 | 4.46 | 4.59 | 4.42 | 4.11 | 45
| 1 | 6 | 4.5 | 4.9554 | 4.38 | 6.76 | 4.53 | 4.6903 | 4.61 | 4.48 | 4.36 | 4.39 | 4.24 | 4.56 | 4.386 | 4.6754 | 4.72 | 4.23 | 4.63 | 4.1301 | 4.3 | 4.08 | 4.26 | 4.56 | 4.73 | 4.4 | 4.13
| Mean |  | 4.4817 | 4.971 | 4.3944 | 6.7217 | 4.5067 | 4.6493 | 4.6183 | 4.4283 | 4.3617 | 4.4283 | 4.1483 | 4.595 | 4.3737 | 4.5812 | 4.76 | 4.1983 | 4.56 | 4.1876 | 4.2667 | 4.1 | 4.265 | 4.5017 | 4.6617 | 4.44 | 4.105
| Median |  | 4.495 | 4.968 | 4.39 | 6.69 | 4.515 | 4.6313 | 4.615 | 4.405 | 4.36 | 4.41 | 4.15 | 4.605 | 4.3575 | 4.5816 | 4.755 | 4.21 | 4.565 | 4.1778 | 4.3 | 4.105 | 4.26 | 4.5 | 4.66 | 4.415 | 4.11
| Std Dev. |  | 0.036009 | 0.014135 | 0.048688 | 0.14878 | 0.052409 | 0.076477 | 0.039707 | 0.073052 | 0.041673 | 0.047504 | 0.068823 | 0.033912 | 0.064447 | 0.07177 | 0.086718 | 0.04535 | 0.0498 | 0.061052 | 0.13663 | 0.063875 | 0.012247 | 0.041191 | 0.05947 | 0.053292 | 0.029496
| Rel.Std.Dev. |  | 0.8035% | 0.2843% | 1.108% | 2.213% | 1.163% | 1.645% | 0.8598% | 1.65% | 0.9554% | 1.073% | 1.659% | 0.738% | 1.474% | 1.567% | 1.822% | 1.08% | 1.092% | 1.458% | 3.202% | 1.558% | 0.2872% | 0.915% | 1.276% | 1.2% | 0.7185%
| PDM3 |  | 0.8873% | 11.9% | -1.077% | 51.31% | 1.45% | 4.66% | 3.964% | -0.3133% | -1.814% | -0.3133% | -6.616% | 3.439% | -1.544% | 3.129% | 7.153% | -5.491% | 2.651% | -5.733% | -3.953% | -7.704% | -3.99% | 1.337% | 4.939% | -0.0507% | -7.592%
| Z-Score (Absolute) |  | 0.16 | 2.07 | 0.18 | 8.88 | 0.26 | 0.81 | 0.69 | 0.05 | 0.3 | 0.05 | 1.13 | 0.6 | 0.26 | 0.55 | 1.24 | 0.94 | 0.47 | 0.98 | 0.67 | 1.32 | 0.68 | 0.24 | 0.86 | 0 | 1.3
| Analytical results for Sb in OREAS 313 (Certified Value 44.4 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 44.3 | 43.57 | 44.064 | 46.14 | 45.03 | 46.6359 | 46.9 | 43.5 | 47 | 45.9 | 38.8 | 43.4 | 38.06 | 42.5661 | 48.5 | 47.15 | 43.14 | 46.97 | 46.4 | 35.35 | 11.7 | 38.51 | 38.52 | 48.4 | 26.6 | 1
| 1 | 2 | 45.3 | 44.63 | 43.848 | 46.11 | 43.2 | 48.0009 | 46.5 | 43.3 | 46.9 | 47.2 | 40.6 | 42.6 | 37.85 | 44.0505 | 56 | 47.14 | 42.9 | 47.59 | 44.8 | 36.8 | 12.9 | 40.9 | 35.53 | 49 | 27.9 | 21
| 1 | 3 | 45.1 | 46.39 | 44.928 | 45.74 | 44.14 | 46.4106 | 45.8 | 44.1 | 46.4 | 47.3 | 40.6 | 39.9 | 38.8 | 41.2006 | 57.9 | 47.24 | 42.63 | 43.87 | 44.4 | 32.75 | 14.3 | 40.47 | 30.25 | 48.2 | 27.3 | 16
| 1 | 4 | 45.9 | 39.57 | 44.5 | 46.92 | 45.85 | 47.4104 | 44.8 | 43.8 | 46.5 | 46.5 | 39.9 | 44 | 38.54 | 40.6606 | 57.7 | 46.5 | 41.7 | 45.4 | 45.6 | 35.06 | 16.8 | 39.23 | 38.81 | 48.9 | 28.2 | 44.4058
| 1 | 5 | 44.8 | 45.7 | 43.89 | 45.99 | 43.84 | 46.0584 | 46.8 | 46.9 | 46.9 | 45.7 | 40.8 | 41.9 | 39.16 | 45.9132 | 52.3 | 47.17 | 42.68 | 47.46 | 41.7 | 32.89 | 11.6 | 39.48 | 32.56 | 49 | 25.2 | 46
| 1 | 6 | 44.5 | 45.3 | 43.416 | 45.78 | 45.13 | 47.1024 | 47.8 | 43.9 | 48.1 | 44.2 | 40.2 | 36.5 | 40.75 | 44.9084 | 57.7 | 47.15 | 42.05 | 45.62 | 47.3 | 33.22 | 11.4 | 41 | 26.88 | 47.5 | 27.2
| Mean |  | 44.9833 | 44.1933 | 44.1077 | 46.1133 | 44.5317 | 46.9364 | 46.4333 | 44.25 | 46.9667 | 46.1333 | 40.15 | 41.3833 | 38.86 | 43.2166 | 55.0167 | 47.0583 | 42.5167 | 46.1517 | 45.0333 | 34.345 | 13.1167 | 39.9317 | 33.7583 | 48.5 | 27.0667
| Median |  | 44.95 | 44.965 | 43.977 | 46.05 | 44.585 | 46.8691 | 46.65 | 43.85 | 46.9 | 46.2 | 40.4 | 42.25 | 38.67 | 43.3083 | 56.85 | 47.15 | 42.655 | 46.295 | 45.2 | 34.14 | 12.3 | 39.975 | 34.045 | 48.65 | 27.25
| Std Dev. |  | 0.58109 | 2.4607 | 0.53329 | 0.42828 | 0.97463 | 0.71035 | 1.0289 | 1.3293 | 0.60553 | 1.1501 | 0.73689 | 2.7824 | 1.0419 | 2.0904 | 3.8348 | 0.27593 | 0.54032 | 1.4494 | 1.9439 | 1.6416 | 2.1104 | 1.0086 | 4.741 | 0.5933 | 1.0727
| Rel.Std.Dev. |  | 1.292% | 5.568% | 1.209% | 0.9288% | 2.189% | 1.513% | 2.216% | 3.004% | 1.289% | 2.493% | 1.835% | 6.723% | 2.681% | 4.837% | 6.97% | 0.5864% | 1.271% | 3.141% | 4.317% | 4.78% | 16.09% | 2.526% | 14.04% | 1.223% | 3.963%
| PDM3 |  | 1.301% | -0.4784% | -0.6713% | 3.845% | 0.2835% | 5.699% | 4.566% | -0.3508% | 5.767% | 3.89% | -9.584% | -6.806% | -12.49% | -2.678% | 23.9% | 5.973% | -4.254% | 3.932% | 1.413% | -22.66% | -70.46% | -10.08% | -23.98% | 9.22% | -39.05%
| Z-Score (Absolute) |  | 0.18 | 0.01 | 0.04 | 0.47 | 0.07 | 0.67 | 0.55 | 0 | 0.68 | 0.47 | 1.03 | 0.72 | 1.35 | 0.26 | 2.7 | 0.7 | 0.44 | 0.48 | 0.2 | 2.49 | 7.81 | 1.08 | 2.63 | 1.07 | 4.31
| Analytical results for Sc in OREAS 313 (Certified Value 8.7 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 8 | 8.6381 | 9 | 8.3 | 9.1 | 9.0463 | 8.4 | 8.4 | 9.1 | 9.7 | 7.8 | 9 | 7.2 | 8.2342 | 7.8 | 8.6 | 8.7 | 9 | 7.4 | 8.5 | 8.8 | 9.2 | 9.6 | 9.5 |  | 1
| 1 | 2 | 8 | 8.6162 | 9 | 8.4 | 8.5 | 9.2523 | 8.3 | 8.3 | 9.2 | 8 | 8 | 9 | 7.4 | 8.4919 | 8.5 | 9 | 8.7 | 9 | 7.6 | 8.8 | 8.3 | 9.1 | 9.7 | 9.5 |  | 22
| 1 | 3 | 8 | 8.6673 | 9 | 8.5 | 8.5 | 8.7565 | 8.3 | 8.4 | 8.7 | 9.7 | 8.1 | 9 | 7.5 | 8.2018 | 8.3 | 9.1 | 8.4 | 9 | 7.2 | 8.7 | 9 | 8.9 | 9.7 | 9.5 |  | 16
| 1 | 4 | 8 | 8.6809 | 9 | 8.4 | 9.4 | 9.2734 | 8.2 | 8.1 | 9 | 9.9 | 8 | 9 | 7.3 | 8.5357 | 7.8 | 8.9 | 8.4 | 9 | 7.4 | 8.7 | 7.9 | 8.7 | 9.8 | 9.6 |  | 8.7012
| 1 | 5 | 8 | 8.6697 | 9 | 8.3 | 8.4 | 8.6573 | 8.4 | 8.1 | 9.5 | 9.6 | 8.2 | 9 | 7.5 | 8.6787 | 9 | 9.1 | 8.6 | 8 | 7.3 | 8.8 | 8.1 | 8.7 | 9.7 | 9.6 |  | 47
| 1 | 6 | 9 | 8.6869 | 9 | 8.5 | 9 | 9.2796 | 8.4 | 8.3 | 9.4 | 9.4 | 8.1 | 9 | 7.7 | 8.3726 | 8.3 | 9.1 | 8.6 | 9 | 7.1 | 8.6 | 8.4 | 9 | 9.6 | 9.3
| Mean |  | 8.1667 | 8.6599 | 9 | 8.4 | 8.8167 | 9.0442 | 8.3333 | 8.2667 | 9.15 | 9.3833 | 8.0333 | 9 | 7.4333 | 8.4192 | 8.2833 | 8.9667 | 8.5667 | 8.8333 | 7.3333 | 8.6833 | 8.4167 | 8.9333 | 9.6833 | 9.5
| Median |  | 8 | 8.6685 | 9 | 8.4 | 8.75 | 9.1493 | 8.35 | 8.3 | 9.15 | 9.65 | 8.05 | 9 | 7.45 | 8.4322 | 8.3 | 9.05 | 8.6 | 9 | 7.35 | 8.7 | 8.35 | 8.95 | 9.7 | 9.5
| Std Dev. |  | 0.40825 | 0.027224 | 0 | 0.089443 | 0.40702 | 0.27703 | 0.08165 | 0.13663 | 0.2881 | 0.6969 | 0.13663 | 0 | 0.17512 | 0.18431 | 0.4535 | 0.19664 | 0.13663 | 0.40825 | 0.17512 | 0.1169 | 0.41673 | 0.20656 | 0.075277 | 0.10954
| Rel.Std.Dev. |  | 4.999% | 0.3144% | 0% | 1.065% | 4.617% | 3.063% | 0.9798% | 1.653% | 3.149% | 7.427% | 1.701% | 0% | 2.356% | 2.189% | 5.475% | 2.193% | 1.595% | 4.622% | 2.388% | 1.346% | 4.951% | 2.312% | 0.7774% | 1.153%
| PDM3 |  | -6.143% | -0.4755% | 3.434% | -3.462% | 1.327% | 3.942% | -4.228% | -4.994% | 5.158% | 7.839% | -7.676% | 3.434% | -14.57% | -3.242% | -4.803% | 3.051% | -1.546% | 1.518% | -15.72% | -0.2056% | -3.27% | 2.668% | 11.29% | 9.18%
| Z-Score (Absolute) |  | NA | 0.09 | NA | 0.43 | 0.41 | 0.86 | 0.56 | 0.69 | 1.07 | 1.54 | 1.16 | NA | 2.35 | 0.39 | 0.66 | 0.7 | 0.09 | NA | 2.55 | 0.14 | 0.39 | 0.64 | 2.13 | 1.77
Results from laboratories 2.01, 2.03, 2.12 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 313 (Indicative Value 2.65 ppm) | [26 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 2 | 2.5084 | 3 | < 5 | 1.8 | < 1.5 | 3 | 3 | 4 | 3 | 5 | 3.4 | 2 | < 2 | 3 | 1.6 | 2.1 | 1.91 | 3 | < 5 | 3 | 2.2 | 2 |  |  | 1
| 1 | 2 | 2 | 2.4736 | 3 | < 5 | 1.7 | < 1.5 | 3 | 3 | 4 | 2 | 4 | 3.3 | 2 | < 2 | 2 | 1.7 | 2.7 | 1.96 | 3 | < 5 | 3 | 1.5 | 2 |  |  | 2
| 1 | 3 | 2 | 2.49 | 3 | < 5 | 1.5 | < 1.5 | 4 | 3 | 4 | 3 | 4 | 2.7 | 2 | < 2 | 2 | 1.7 | 3.3 | 1.81 | 3 | < 5 | 3 | 2.1 | 2 |  |  | 16
| 1 | 4 | 2 | 2.4487 | 3 | < 5 | 1.8 | < 1.5 | 3 | 3 | 4 | 3 | 3 | 2.7 | 2 | < 2 | 2 | 1.6 | 1.8 | 1.81 | 3 | < 5 | 4 | 1.8 | 2 |  |  | 2.6471
| 1 | 5 | 2 | 2.4716 | 3 | < 5 | 1.9 | < 1.5 | 4 | 3 | 4 | 3 | 4 | 2.9 | 2 | < 2 | 2 | 1.7 | 2.8 | 1.88 | 3 | < 5 | 4 | 1.4 | 2 |  |  | 8
| 1 | 6 | 2 | 2.5532 | 3 | < 5 | 1.9 | < 1.5 | 4 | 4 | 4 | 3 | 5 | 3.1 | 2 | < 2 | 2 | 1.6 | 1.6 | 1.85 | 3 | < 5 | 4 | 1.7 | 2
| Mean |  | 2 | 2.4909 | 3 |  | 1.7667 |  | 3.5 | 3.1667 | 4 | 2.8333 | 4.1667 | 3.0167 | 2 |  | 2.1667 | 1.65 | 2.3833 | 1.87 | 3 |  | 3.5 | 1.7833 | 2
| Median |  | 2 | 2.4818 | 3 |  | 1.8 |  | 3.5 | 3 | 4 | 3 | 4 | 3 | 2 |  | 2 | 1.65 | 2.4 | 1.865 | 3 |  | 3.5 | 1.75 | 2
| Std Dev. |  | 0 | 0.036444 | 0 |  | 0.15055 |  | 0.54772 | 0.40825 | 0 | 0.40825 | 0.75277 | 0.29944 | 0 |  | 0.40825 | 0.054772 | 0.65549 | 0.058992 | 0 |  | 0.54772 | 0.31885 | 0
| Rel.Std.Dev. |  | 0% | 1.463% | 0% |  | 8.522% |  | 15.65% | 12.89% | 0% | 14.41% | 18.07% | 9.926% | 0% |  | 18.84% | 3.32% | 27.5% | 3.155% | 0% |  | 15.65% | 17.88% | 0%
| PDM3 |  | -24.44% | -5.899% | 13.33% |  | -33.26% |  | 32.22% | 19.63% | 51.11% | 7.037% | 57.41% | 13.96% | -24.44% |  | -18.15% | -37.67% | -9.963% | -29.36% | 13.33% |  | 32.22% | -32.63% | -24.44%
| Z-Score (Absolute) |  | 0.63 | 0 | 0.65 | 0.01 | 0.93 | 2.23 | 1.29 | 0.87 | 1.94 | 0.44 | 2.15 | 0.67 | 0.63 | 1.91 | 0.42 | 1.08 | 0.14 | 0.8 | 0.65 | 0.01 | 1.29 | 0.91 | 0.63
| Analytical results for Sm in OREAS 313 (Certified Value 4.87 ppm) | [26 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
3
| 1 | 1 | 4.3859 | 4.88 | 5.137 | 5.2 | 5.03 | 5 | 4.19 | [18 empty] | 1
| 1 | 2 | 4.3778 | 4.86 | 5.3783 | 5.2 | 4.83 | 5.1 | 4.02 | [18 empty] | 24
| 1 | 3 | 4.4138 | 4.86 | 5.2698 | 5 | 5 | 5.1 | 4.03 | [18 empty] | 16
| 1 | 4 | 4.4165 | 4.76 | 5.4176 | 4.3 | 4.78 | 4.9 | 4.29 | [18 empty] | 4.8749
| 1 | 5 | 4.3811 | 4.88 | 5.1496 | 4.6 | 5.01 | 5 | 4.19 | [18 empty] | 48
| 1 | 6 | 4.4647 | 4.78 | 5.105 | 4.7 | 4.93 | 4.9 | 4.11
| Mean |  | 4.4066 | 4.8367 | 5.2429 | 4.8333 | 4.93 | 5 | 4.1383
| Median |  | 4.3999 | 4.86 | 5.2097 | 4.85 | 4.965 | 5 | 4.15
| Std Dev. |  | 0.032966 | 0.052789 | 0.13305 | 0.36148 | 0.10373 | 0.089443 | 0.10477
| Rel.Std.Dev. |  | 0.7481% | 1.091% | 2.538% | 7.479% | 2.104% | 1.789% | 2.532%
| PDM3 |  | -9.606% | -0.7848% | 7.549% | -0.8532% | 1.13% | 2.566% | -15.11%
| Z-Score (Absolute) |  | 1.78 | 0 | 1.68 | 0.01 | 0.39 | 0.67 | 2.88
| Analytical results for Sn in OREAS 313 (Certified Value 2.59 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.4 | 3.8 | < 5 | 2.6 | 2.6404 | 2.9 | 2.6 | 2.9 | 2.4 | 2.3 | 3 | 2.2 | 2.6099 | 2 | 2.7 | 2.7 | 2.6 | 2.9 | 2.2 | 2.42 | 2.7 | 2.7 | 6.2 |  |  | 1
| 1 | 2 | 2.4 | 3.9 | < 5 | 2.6 | 2.7764 | 2.9 | 2.6 | 2.9 | 2.5 | 2.5 | 3 | 2.2 | 2.6764 | 1.8 | 2.7 | 2.6 | 2.5 | 2.7 | 2.1 | 2.44 | 2.6 | 2.9 | 6.2 |  |  | 31
| 1 | 3 | 2.4 | 3.6 | < 5 | 2.7 | 2.6984 | 2.9 | 2.7 | 2.9 | 2.4 | 2.5 | 3 | 2.2 | 2.5955 | 2.2 | 2.7 | 2.6 | 2.5 | 2.7 | 2.1 | 2.67 | 2.8 | 2.7 | 6.2 |  |  | 16
| 1 | 4 | 2.4 | 3.6 | < 5 | 2.7 | 2.7902 | 2.8 | 2.6 | 2.8 | 2.5 | 2.5 | 3 | 2.2 | 2.3845 | 1.9 | 2.7 | 2.6 | 2.6 | 2.8 | 2.1 | 2.43 | 2.8 | 2.7 | 6.2 |  |  | 2.5852
| 1 | 5 | 2.5 | 3.8 | < 5 | 2.6 | 2.6404 | 2.9 | 2.5 | 2.9 | 2.3 | 2.5 | 3 | 2.2 | 2.657 | 1.9 | 2.7 | 2.7 | 2.4 | 2.8 | 2.2 | 2.38 | 2.7 | 2.8 | 6.2 |  |  | 49
| 1 | 6 | 2.5 | 3.8 | < 5 | 2.6 | 2.7025 | 3 | 2.6 | 3 | 2.3 | 2.5 | 3 | 2.2 | 2.3633 | 2 | 2.7 | 2.6 | 2.6 | 2.7 | 2.1 | 2.63 | 2.7 | 2.8 | 6.2
| Mean |  | 2.4333 | 3.75 |  | 2.6333 | 2.708 | 2.9 | 2.6 | 2.9 | 2.4 | 2.4667 | 3 | 2.2 | 2.5478 | 1.9667 | 2.7 | 2.6333 | 2.5333 | 2.7667 | 2.1333 | 2.495 | 2.7167 | 2.7667 | 6.2
| Median |  | 2.4 | 3.8 |  | 2.6 | 2.7005 | 2.9 | 2.6 | 2.9 | 2.4 | 2.5 | 3 | 2.2 | 2.6027 | 1.95 | 2.7 | 2.6 | 2.55 | 2.75 | 2.1 | 2.435 | 2.7 | 2.75 | 6.2
| Std Dev. |  | 0.05164 | 0.12247 |  | 0.05164 | 0.06436 | 0.063246 | 0.063246 | 0.063246 | 0.089443 | 0.08165 | 0 | 4.8648e-16 | 0.13806 | 0.13663 | 4.8648e-16 | 0.05164 | 0.08165 | 0.08165 | 0.05164 | 0.12243 | 0.075277 | 0.08165 | 0
| Rel.Std.Dev. |  | 2.122% | 3.266% |  | 1.961% | 2.377% | 2.181% | 2.433% | 2.181% | 3.727% | 3.31% | 0% | 2.211e-14% | 5.419% | 6.947% | 1.802e-14% | 1.961% | 3.223% | 2.951% | 2.421% | 4.907% | 2.771% | 2.951% | 0%
| PDM3 |  | -5.876% | 45.05% |  | 1.861% | 4.751% | 12.18% | 0.5713% | 12.18% | -7.165% | -4.586% | 16.04% | -14.9% | -1.449% | -23.93% | 4.439% | 1.861% | -2.007% | 7.018% | -17.48% | -3.49% | 5.084% | 7.018% | 139.8%
| Z-Score (Absolute) |  | 0.82 | 5.09 | 0.52 | 0.07 | 0.41 | 1.27 | 0.07 | 1.27 | 0.97 | 0.67 | NA | 1.87 | 0.31 | 2.92 | 0.37 | 0.07 | 0.37 | 0.67 | 2.17 | 0.55 | 0.45 | 0.67 | 16.11
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 313 (Certified Value 41.7 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 39.9 | 41.7 | 41 | 41.44 | 41.06 | 44.8045 | 42.4 | 41 | 42.6 | 38.9 | 37.8 | 39.1 | 39.6 | 44.553 | 26.1 | 42.25 | 42 | 41 | 33.5 | 42.8 | 39.5 | 42.89 | 44.5 | 43 |  | 1
| 1 | 2 | 40 | 39.86 | 40 | 41.58 | 38.82 | 45.0808 | 42 | 40.9 | 41.6 | 42.3 | 37.8 | 35.6 | 39.3 | 43.5858 | 32.2 | 42.7 | 42 | 40 | 33.1 | 42.9 | 40.1 | 42.64 | 45.9 | 42 |  | 32
| 1 | 3 | 39.6 | 41.83 | 40 | 42.72 | 40.17 | 43.3465 | 42 | 42.9 | 42.8 | 39.5 | 39.5 | 27.9 | 41 | 44.2987 | 29.9 | 42.72 | 41 | 40 | 34.6 | 42.2 | 40.3 | 41.76 | 45.2 | 43 |  | 16
| 1 | 4 | 40.4 | 39.38 | 41 | 42.17 | 42.59 | 44.8321 | 42.4 | 41.4 | 41.4 | 42.5 | 39.3 | 30.2 | 39.7 | 45.4588 | 30.4 | 42.88 | 40 | 40 | 32.1 | 43 | 39.3 | 41.39 | 45.8 | 44 |  | 41.676
| 1 | 5 | 39.7 | 42.39 | 40 | 41.77 | 39.5 | 42.9564 | 42.5 | 37 | 41.7 | 38.7 | 38.8 | 34.9 | 40.6 | 45.714 | 30.5 | 43.63 | 41 | 37 | 32.8 | 43.3 | 39.5 | 41.71 | 46.2 | 44 |  | 50
| 1 | 6 | 39.1 | 40.7 | 42 | 42.16 | 40.57 | 44.2188 | 43.2 | 41.1 | 43.3 | 38.3 | 39.8 | 34.1 | 39.3 | 43.8922 | 32.1 | 42.67 | 41 | 41 | 32.8 | 42.8 | 39.5 | 41.73 | 45.8 | 43
| Mean |  | 39.7833 | 40.9767 | 40.6667 | 41.9733 | 40.4517 | 44.2065 | 42.4167 | 40.7167 | 42.2333 | 40.0333 | 38.8333 | 33.6333 | 39.9167 | 44.5838 | 30.2 | 42.8083 | 41.1667 | 39.8333 | 33.15 | 42.8333 | 39.7 | 42.02 | 45.5667 | 43.1667
| Median |  | 39.8 | 41.2 | 40.5 | 41.965 | 40.37 | 44.5116 | 42.4 | 41.05 | 42.15 | 39.2 | 39.05 | 34.5 | 39.65 | 44.4258 | 30.45 | 42.71 | 41 | 40 | 32.95 | 42.85 | 39.5 | 41.745 | 45.8 | 43
| Std Dev. |  | 0.43551 | 1.1935 | 0.8165 | 0.47149 | 1.3118 | 0.87359 | 0.44008 | 1.9651 | 0.77115 | 1.8747 | 0.8641 | 4.0058 | 0.71391 | 0.84847 | 2.2199 | 0.45402 | 0.75277 | 1.472 | 0.84558 | 0.36148 | 0.4 | 0.59766 | 0.61536 | 0.75277
| Rel.Std.Dev. |  | 1.095% | 2.913% | 2.008% | 1.123% | 3.243% | 1.976% | 1.038% | 4.826% | 1.826% | 4.683% | 2.225% | 11.91% | 1.788% | 1.903% | 7.351% | 1.061% | 1.829% | 3.695% | 2.551% | 0.8439% | 1.008% | 1.422% | 1.35% | 1.744%
| PDM3 |  | -4.541% | -1.678% | -2.422% | 0.7133% | -2.938% | 6.072% | 1.777% | -2.302% | 1.337% | -3.942% | -6.821% | -19.3% | -4.222% | 6.977% | -27.54% | 2.717% | -1.222% | -4.422% | -20.46% | 2.777% | -4.741% | 0.8253% | 9.335% | 3.577%
| Z-Score (Absolute) |  | 0.56 | 0.07 | 0.09 | 0.59 | 0.21 | 1.77 | 0.83 | 0.07 | 0.73 | 0.43 | 1.06 | 3.8 | 0.49 | 1.97 | 5.61 | 1.03 | 0.17 | 0.53 | 4.05 | 1.05 | 0.6 | 0.62 | 2.49 | 1.22
| Analytical results for Ta in OREAS 313 (Certified Value 0.24 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.23 | 0.54136 | 0.3 | 0.68 | 0.27 | 0.36514 | 0.29 | 0.31 | 0.22 | 0.22 | 0.23 | 0.6 | 0.15 | 0.221 | 0.31 | 0.37 | 0.2 | 0.2 | 0.12 | 0.19 | 0.21 | 0.39 | 0.2 |  |  | 1
| 1 | 2 | 0.22 | 0.48995 | 0.4 | 0.75 | 0.21 | 0.33636 | 0.29 | 0.29 | 0.24 | 0.25 | 0.25 | 0.5 | 0.18 | 0.2417 | 0.14 | 0.27 | 0.2 | 0.2 | 0.1 | 0.18 | 0.19 | 0.4 | 0.2 |  |  | 11
| 1 | 3 | 0.22 | 0.53653 | 0.3 | 0.74 | 0.26 | 0.38758 | 0.29 | 0.3 | 0.23 | 0.27 | 0.25 | 0.5 | 0.17 | 0.2331 | 0.33 | 0.28 | 0.2 | 0.2 | 0.12 | 0.2 | 0.26 | 0.38 | 0.2 |  |  | 16
| 1 | 4 | 0.23 | 0.49131 | 0.4 | 0.71 | 0.27 | 0.35094 | 0.29 | 0.3 | 0.22 | 0.24 | 0.26 | 0.5 | 0.15 | 0.2279 | 0.14 | 0.39 | 0.2 | 0.2 | 0.12 | 0.2 | 0.22 | 0.4 | 0.2 |  |  | 0.2391
| 1 | 5 | 0.22 | 0.47682 | 0.3 | 0.63 | 0.21 | 0.36529 | 0.28 | 0.29 | 0.15 | 0.25 | 0.24 | 0.5 | 0.19 | 0.2363 | 0.15 | 0.32 | 0.2 | 0.2 | 0.11 | 0.16 | 0.21 | 0.4 | 0.2 |  |  | 51
| 1 | 6 | 0.22 | 0.49614 | 0.3 | 0.7 | 0.24 | 0.31514 | 0.3 | 0.3 | 0.24 | 0.24 | 0.26 | 0.4 | 0.17 | 0.219 | 0.19 | 0.36 | 0.2 | 0.2 | 0.12 | 0.19 | 0.19 | 0.42 | 0.2
| Mean |  | 0.22333 | 0.50535 | 0.33333 | 0.70167 | 0.24333 | 0.35341 | 0.29 | 0.29833 | 0.21667 | 0.245 | 0.24833 | 0.5 | 0.16833 | 0.22983 | 0.21 | 0.33167 | 0.2 | 0.2 | 0.115 | 0.18667 | 0.21333 | 0.39833 | 0.2
| Median |  | 0.22 | 0.49372 | 0.3 | 0.705 | 0.25 | 0.35804 | 0.29 | 0.3 | 0.225 | 0.245 | 0.25 | 0.5 | 0.17 | 0.2305 | 0.17 | 0.34 | 0.2 | 0.2 | 0.12 | 0.19 | 0.21 | 0.4 | 0.2
| Std Dev. |  | 0.005164 | 0.02684 | 0.05164 | 0.043551 | 0.028048 | 0.025329 | 0.0063246 | 0.0075277 | 0.033862 | 0.016432 | 0.01169 | 0.063246 | 0.016021 | 0.0088593 | 0.087407 | 0.049565 | 3.0405e-17 | 3.0405e-17 | 0.0083666 | 0.015055 | 0.02582 | 0.013292 | 3.0405e-17
| Rel.Std.Dev. |  | 2.312% | 5.311% | 15.49% | 6.207% | 11.53% | 7.167% | 2.181% | 2.523% | 15.63% | 6.707% | 4.708% | 12.65% | 9.517% | 3.855% | 41.62% | 14.94% | 1.52e-14% | 1.52e-14% | 7.275% | 8.065% | 12.1% | 3.337% | 1.52e-14%
| PDM3 |  | -6.596% | 111.4% | 39.41% | 193.5% | 1.768% | 47.8% | 21.29% | 24.77% | -9.384% | 2.466% | 3.86% | 109.1% | -29.6% | -3.878% | -12.17% | 38.71% | -16.35% | -16.35% | -51.9% | -21.93% | -10.78% | 66.59% | -16.35%
| Z-Score (Absolute) |  | 0.28 | 3.52 | NA | 6.17 | 0.01 | 1.47 | 0.62 | 0.73 | 0.37 | 0.01 | 0.06 | NA | 1.02 | 0.19 | 0.46 | 1.18 | NA | NA | 1.74 | 0.78 | 0.42 | 2.08 | NA
Results from laboratories 2.03, 2.12, 2.17, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 313 (Certified Value 0.51 ppm) | [26 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.13 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [14 empty] | 2
2
| 1 | 1 | 0.4178 | 0.484 | 0.62 | 0.56498 | 0.6 | 0.41 | 0.55 | 0.5 | 0.39 | 0.53 | 0.51 | [14 empty] | 1
| 1 | 2 | 0.3416 | 0.496 | 0.62 | 0.56532 | 0.6 | 0.41 | 0.54 | 0.4 | 0.36 | 0.53 | 0.55 | [14 empty] | 12
| 1 | 3 | 0.4216 | 0.485 | 0.65 | 0.52911 | 0.6 | 0.43 | 0.55 | 0.5 | 0.36 | 0.49 | 0.55 | [14 empty] | 16
| 1 | 4 | 0.4034 | 0.481 | 0.63 | 0.54023 | 0.6 | 0.43 | 0.51 | 0.5 | 0.36 | 0.49 | 0.5 | [14 empty] | 0.50888
| 1 | 5 | 0.4042 | 0.48 | 0.61 | 0.49835 | 0.6 | 0.44 | 0.56 | 0.5 | 0.37 | 0.53 | 0.51 | [14 empty] | 52
| 1 | 6 | 0.4178 | 0.467 | 0.61 | 0.51764 | 0.6 | 0.43 | 0.55 | 0.5 | 0.35 | 0.51 | 0.59
| Mean |  | 0.40107 | 0.48217 | 0.62333 | 0.53594 | 0.6 | 0.425 | 0.54333 | 0.48333 | 0.365 | 0.51333 | 0.535
| Median |  | 0.411 | 0.4825 | 0.62 | 0.53467 | 0.6 | 0.43 | 0.55 | 0.5 | 0.36 | 0.52 | 0.53
| Std Dev. |  | 0.03011 | 0.0093684 | 0.015055 | 0.026533 | 0 | 0.012247 | 0.017512 | 0.040825 | 0.013784 | 0.019664 | 0.034496
| Rel.Std.Dev. |  | 7.508% | 1.943% | 2.415% | 4.951% | 0% | 2.882% | 3.223% | 8.447% | 3.776% | 3.831% | 6.448%
| PDM3 |  | -21.19% | -5.25% | 22.49% | 5.317% | 17.91% | -16.48% | 6.77% | -5.021% | -28.27% | 0.8745% | 5.132%
| Z-Score (Absolute) |  | 2.43 | 0.67 | 2.38 | 0.49 | NA | 1.91 | 0.65 | NA | 3.21 | 0 | 0.47
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 313 (Certified Value 0.084 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.1 | < 0.5 | 2.75 | < 0.2 | 0.08 | 0.06 | 0.09 | 0.08 | 0.07 | 0.2 | 0.08 | 0.1223 | 0.34 | 0.07 | < 0.03 | < 0.5 | 7.17 | < 0.05 | 0.12 | < 0.1 | [5 empty] | 1
| 1 | 2 | 0.1 | < 0.5 | 2.64 | < 0.2 | 0.07 | < 0.05 | 0.09 | 0.09 | 0.07 | < 0.1 | 0.08 | 0.1146 | 0.34 | 0.08 | 0.07 | < 0.5 | 6.98 | < 0.05 | 0.12 | < 0.1 | [5 empty] | 13
| 1 | 3 | 0.1 | < 0.5 | 2.67 | < 0.2 | 0.08 | 0.07 | 0.08 | 0.08 | 0.07 | < 0.1 | 0.09 | 0.1164 | 0.55 | 0.14 | < 0.03 | < 0.5 | 7.56 | < 0.05 | 0.11 | < 0.1 | [5 empty] | 16
| 1 | 4 | 0.1 | < 0.5 | 2.81 | < 0.2 | 0.09 | 0.07 | 0.09 | 0.07 | < 0.05 | < 0.1 | 0.06 | 0.1058 | 0.02 | 0.06 | < 0.03 | < 0.5 | 6.88 | < 0.05 | 0.13 | < 0.1 | [5 empty] | 0.084432
| 1 | 5 | 0.1 | < 0.5 | 2.86 | < 0.2 | 0.11 | 0.11 | 0.08 | 0.06 | < 0.05 | < 0.1 | 0.07 | 0.1214 | 0.19 | 0.07 | < 0.03 | < 0.5 | 7.06 | < 0.05 | 0.11 | < 0.1 | [5 empty] | 53
| 1 | 6 | 0.1 | < 0.5 | 2.61 | < 0.2 | 0.07 | 0.09 | 0.1 | 0.07 | < 0.05 | 0.1 | 0.07 | 0.1054 | 0.33 | 0.07 | < 0.03 | < 0.5 | 7.19 | < 0.05 | 0.12 | < 0.1
| Mean |  | 0.1 |  | 2.7233 |  | 0.083333 | 0.08 | 0.088333 | 0.075 | 0.07 | 0.15 | 0.075 | 0.11432 | 0.295 | 0.081667 | 0.07 |  | 7.14 |  | 0.11833
| Median |  | 0.1 |  | 2.71 |  | 0.08 | 0.07 | 0.09 | 0.075 | 0.07 | 0.15 | 0.075 | 0.1155 | 0.335 | 0.07 | 0.07 |  | 7.115 |  | 0.12
| Std Dev. |  | 1.5202e-17 |  | 0.099532 |  | 0.015055 | 0.02 | 0.0075277 | 0.010488 | 0 | 0.070711 | 0.010488 | 0.0073535 | 0.17717 | 0.029269 |  |  | 0.23639 |  | 0.0075277
| Rel.Std.Dev. |  | 1.52e-14% |  | 3.655% |  | 18.07% | 25% | 8.522% | 13.98% | 0% | 47.14% | 13.98% | 6.433% | 60.06% | 35.84% |  |  | 3.311% |  | 6.361%
| PDM3 |  | 18.44% |  | 3125% |  | -1.301% | -5.249% | 4.621% | -11.17% | -17.09% | 77.66% | -11.17% | 35.4% | 249.4% | -3.275% | -17.09% |  | 8357% |  | 40.15%
| Z-Score (Absolute) |  | NA | NA | 53.41 | 0.34 | 0 | 0.25 | 0.1 | 0.17 | 0.72 | NA | 0.17 | 0.63 | 4.28 | 0.03 | 1.2 | 3.37 | 142.75 | 1.18 | 0.71 | 0.67
Results from laboratories 2.01, 2.03 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 313 (Certified Value 9.98 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 9.47 | 10.139 | 11.66 | 10.14 | 10.4328 | 10.4 | 9.77 | 10.35 | 10.3 | 8.55 | 8.8 | 8.59 | 8.0626 | 9.5 | 10.11 | 8.9 | 9.6 | 9.32 | 9.8 | 10.25 | 10.04 | 10.6 | 10.6 |  |  | 1
| 1 | 2 | 9.69 | 10.174 | 11.64 | 10.4 | 10.7317 | 10.6 | 9.96 | 9.34 | 11.5 | 8.83 | 7.8 | 8.89 | 7.9117 | 9.4 | 10.3 | 9 | 9.3 | 9.05 | 9.8 | 10.19 | 10.43 | 10.9 | 10.9 |  |  | 14
| 1 | 3 | 9.34 | 10.159 | 11.68 | 10.44 | 10.6856 | 10.25 | 9.91 | 9.39 | 9.95 | 8.85 | 7.5 | 8.85 | 7.7409 | 9.7 | 10.42 | 9 | 9.6 | 8.82 | 9.6 | 10.79 | 11.17 | 10.9 | 10.8 |  |  | 16
| 1 | 4 | 9.74 | 10.216 | 11.93 | 10.43 | 10.8767 | 10.5 | 9.45 | 10.5 | 10.5 | 8.48 | 6.6 | 8.54 | 7.7839 | 9.1 | 10.32 | 8.7 | 9.5 | 8.88 | 9.6 | 10.3 | 9.74 | 10.8 | 11 |  |  | 9.9832
| 1 | 5 | 9.68 | 10.141 | 11.74 | 10.43 | 10.5653 | 10.7 | 10 | 10.3 | 10 | 9.2 | 7.8 | 8.92 | 8.1202 | 9.7 | 10.55 | 9 | 8.9 | 8.66 | 10.2 | 10.26 | 9.58 | 11 | 10.3 |  |  | 54
| 1 | 6 | 9.42 | 10.225 | 11.83 | 10.65 | 10.4546 | 10.6 | 9.79 | 9.9 | 10.25 | 8.71 | 7.7 | 9.52 | 7.7175 | 9.2 | 10.45 | 8.6 | 9.7 | 8.82 | 9.8 | 11.03 | 10.38 | 10.8 | 10.3
| Mean |  | 9.5567 | 10.1757 | 11.7467 | 10.415 | 10.6245 | 10.5083 | 9.8133 | 9.9633 | 10.4167 | 8.77 | 7.7 | 8.885 | 7.8895 | 9.4333 | 10.3583 | 8.8667 | 9.4333 | 8.925 | 9.8 | 10.47 | 10.2233 | 10.8333 | 10.65
| Median |  | 9.575 | 10.1665 | 11.71 | 10.43 | 10.6255 | 10.55 | 9.85 | 10.1 | 10.275 | 8.77 | 7.75 | 8.87 | 7.8478 | 9.45 | 10.37 | 8.95 | 9.55 | 8.85 | 9.8 | 10.28 | 10.21 | 10.85 | 10.7
| Std Dev. |  | 0.16717 | 0.037114 | 0.11308 | 0.16257 | 0.17201 | 0.16253 | 0.20007 | 0.50441 | 0.56804 | 0.25729 | 0.70427 | 0.34979 | 0.17117 | 0.25033 | 0.15198 | 0.17512 | 0.29439 | 0.23063 | 0.21909 | 0.35094 | 0.5735 | 0.13663 | 0.30166
| Rel.Std.Dev. |  | 1.749% | 0.3647% | 0.9626% | 1.561% | 1.619% | 1.547% | 2.039% | 5.063% | 5.453% | 2.934% | 9.146% | 3.937% | 2.17% | 2.654% | 1.467% | 1.975% | 3.121% | 2.584% | 2.236% | 3.352% | 5.61% | 1.261% | 2.833%
| PDM3 |  | -4.272% | 1.928% | 17.66% | 4.326% | 6.424% | 5.261% | -1.701% | -0.1986% | 4.342% | -12.15% | -22.87% | -11% | -20.97% | -5.508% | 3.758% | -11.18% | -5.508% | -10.6% | -1.835% | 4.877% | 2.406% | 8.516% | 6.68%
| Z-Score (Absolute) |  | 0.52 | 0.27 | 2.27 | 0.57 | 0.84 | 0.69 | 0.19 | 0 | 0.58 | 1.52 | 2.88 | 1.37 | 2.64 | 0.67 | 0.5 | 1.4 | 0.67 | 1.32 | 0.21 | 0.64 | 0.33 | 1.11 | 0.87
| Analytical results for Ti in OREAS 313 (Certified Value 0.155 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.15 | 0.13802 | 0.15466 | 0.2 | 0.1624 | 0.16147 | 0.176 | 0.175 | 0.143 | 0.168 | 0.163 | 0.239 | 0.142 | 0.1858 | 0.137 | 0.1573 | 0.153 | 0.162 | 0.1 | 0.14 | 0.16 | 0.1499 | 0.19 | 0.154 |  | 1
| 1 | 2 | 0.148 | 0.13718 | 0.15219 | 0.21 | 0.1389 | 0.16538 | 0.17 | 0.175 | 0.149 | 0.157 | 0.161 | 0.238 | 0.142 | 0.1902 | 0.153 | 0.1505 | 0.152 | 0.16 | 0.1 | 0.14 | 0.16 | 0.1477 | 0.19 | 0.148 |  | 37
| 1 | 3 | 0.152 | 0.13822 | 0.15751 | 0.21 | 0.1505 | 0.1575 | 0.172 | 0.18 | 0.15 | 0.167 | 0.161 | 0.239 | 0.147 | 0.1938 | 0.144 | 0.143 | 0.15 | 0.16 | 0.1 | 0.14 | 0.16 | 0.1546 | 0.18 | 0.15 |  | 16
| 1 | 4 | 0.15 | 0.13883 | 0.1486 | 0.21 | 0.1658 | 0.16462 | 0.175 | 0.175 | 0.147 | 0.17 | 0.161 | 0.237 | 0.145 | 0.1966 | 0.13 | 0.1623 | 0.148 | 0.165 | 0.1 | 0.14 | 0.16 | 0.157 | 0.19 | 0.149 |  | 0.15459
| 1 | 5 | 0.152 | 0.13968 | 0.15295 | 0.18 | 0.1442 | 0.15556 | 0.173 | 0.17 | 0.126 | 0.17 | 0.161 | 0.237 | 0.148 | 0.1885 | 0.15 | 0.157 | 0.15 | 0.16 | 0.1 | 0.14 | 0.16 | 0.152 | 0.19 | 0.149 |  | 55
| 1 | 6 | 0.151 | 0.13915 | 0.14839 | 0.18 | 0.1507 | 0.16038 | 0.173 | 0.18 | 0.15 | 0.163 | 0.161 | 0.24 | 0.141 | 0.1947 | 0.148 | 0.1591 | 0.155 | 0.163 | 0.1 | 0.14 | 0.16 | 0.1486 | 0.19 | 0.145
| Mean |  | 0.1505 | 0.13851 | 0.15238 | 0.19833 | 0.15208 | 0.16082 | 0.17317 | 0.17583 | 0.14417 | 0.16583 | 0.16133 | 0.23833 | 0.14417 | 0.1916 | 0.14367 | 0.15487 | 0.15133 | 0.16167 | 0.1 | 0.14 | 0.16 | 0.15163 | 0.18833 | 0.14917
| Median |  | 0.1505 | 0.13853 | 0.15257 | 0.205 | 0.1506 | 0.16092 | 0.173 | 0.175 | 0.148 | 0.1675 | 0.161 | 0.2385 | 0.1435 | 0.192 | 0.146 | 0.15715 | 0.151 | 0.161 | 0.1 | 0.14 | 0.16 | 0.15095 | 0.19 | 0.149
| Std Dev. |  | 0.0015166 | 0.00088904 | 0.0035238 | 0.01472 | 0.010344 | 0.0038592 | 0.002137 | 0.0037639 | 0.0092826 | 0.0050365 | 0.0008165 | 0.0012111 | 0.0029269 | 0.0041149 | 0.0086872 | 0.006979 | 0.0025033 | 0.0020656 | 1.5202e-17 | 0 | 0 | 0.003617 | 0.0040825 | 0.0029269
| Rel.Std.Dev. |  | 1.008% | 0.6418% | 2.312% | 7.422% | 6.801% | 2.4% | 1.234% | 2.141% | 6.439% | 3.037% | 0.5061% | 0.5081% | 2.03% | 2.148% | 6.047% | 4.506% | 1.654% | 1.278% | 1.52e-14% | 0% | 0% | 2.385% | 2.168% | 1.962%
| PDM3 |  | -2.643% | -10.4% | -1.425% | 28.3% | -1.619% | 4.031% | 12.02% | 13.74% | -6.74% | 7.276% | 4.365% | 54.18% | -6.74% | 23.94% | -7.064% | 0.1814% | -2.104% | 4.58% | -35.31% | -9.436% | 3.502% | -1.91% | 21.83% | -3.506%
| Z-Score (Absolute) |  | 0.22 | 1.08 | 0.09 | 3.19 | 0.11 | 0.51 | 1.39 | 1.58 | 0.67 | 0.87 | 0.55 | 6.04 | 0.67 | 2.71 | 0.71 | 0.09 | 0.16 | 0.57 | 3.82 | 0.97 | 0.45 | 0.14 | 2.47 | 0.32
| Analytical results for Tl in OREAS 313 (Certified Value 20.1 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 20.4 | 18.432 | 20.27 | 19.55 | 21.5526 | 20.1 | 19.65 | 19.65 | 19.85 | 16 | 20.2 | 21.33 | 22.2741 | 18.5 | 20.57 | 19.47 | 16.331 | 20.8 | 18 | 18.92 | 20.53 | 24.01 | 13.6 |  |  | 1
| 1 | 2 | 20.3 | 18.72 | 20.3 | 19.25 | 21.6567 | 20.2 | 19.95 | 20.2 | 22.1 | 16.3 | 20.8 | 22.42 | 22.6401 | 18.9 | 20.63 | 19.78 | 16.581 | 20.2 | 18.2 | 19.27 | 20.92 | 23.52 | 14.2 |  |  | 15
| 1 | 3 | 20.5 | 18.24 | 20.21 | 19.89 | 22.0058 | 20 | 19.4 | 19.9 | 19.9 | 16.5 | 19.2 | 22.3 | 22.2967 | 18.8 | 20.61 | 19.27 | 15.519 | 20.7 | 17.4 | 19.43 | 21.12 | 23.4 | 13.9 |  |  | 16
| 1 | 4 | 20.4 | 18.528 | 20.54 | 19.83 | 22.5576 | 19.9 | 18.55 | 19.45 | 20.2 | 16.1 | 19.4 | 21.95 | 22.3821 | 18.2 | 20.43 | 19.19 | 16.017 | 20.2 | 17.9 | 19.41 | 20.9 | 23.79 | 14.4 |  |  | 20.0787
| 1 | 5 | 20 | 18.432 | 20.7 | 20.06 | 22.0672 | 20.3 | 18.15 | 20.4 | 19.55 | 16.7 | 19.9 | 21.85 | 21.9247 | 19.1 | 20.74 | 19.26 | 16.651 | 19.3 | 18.6 | 18.42 | 20.75 | 22.83 | 13.3 |  |  | 56
| 1 | 6 | 20.5 | 18.6 | 20.29 | 20.31 | 22.0813 | 20.3 | 19.6 | 20.5 | 19.75 | 16.55 | 20.4 | 22.1 | 21.8979 | 18.8 | 20.36 | 19.41 | 15.941 | 21.8 | 18 | 18.32 | 21.12 | 22.98 | 13.7
| Mean |  | 20.35 | 18.492 | 20.385 | 19.815 | 21.9869 | 20.1333 | 19.2167 | 20.0167 | 20.225 | 16.3583 | 19.9833 | 21.9917 | 22.2359 | 18.7167 | 20.5567 | 19.3967 | 16.1733 | 20.5 | 18.0167 | 18.9617 | 20.89 | 23.4217 | 13.85
| Median |  | 20.4 | 18.48 | 20.295 | 19.86 | 22.0365 | 20.15 | 19.5 | 20.05 | 19.875 | 16.4 | 20.05 | 22.025 | 22.2854 | 18.8 | 20.59 | 19.34 | 16.174 | 20.45 | 18 | 19.095 | 20.91 | 23.46 | 13.8
| Std Dev. |  | 0.18708 | 0.16471 | 0.19149 | 0.37426 | 0.35751 | 0.1633 | 0.70545 | 0.42032 | 0.94274 | 0.27279 | 0.608 | 0.38727 | 0.2832 | 0.31885 | 0.13909 | 0.21445 | 0.4303 | 0.82946 | 0.392 | 0.49443 | 0.22627 | 0.45543 | 0.40373
| Rel.Std.Dev. |  | 0.9193% | 0.8907% | 0.9394% | 1.889% | 1.626% | 0.8111% | 3.671% | 2.1% | 4.661% | 1.668% | 3.043% | 1.761% | 1.274% | 1.704% | 0.6766% | 1.106% | 2.661% | 4.046% | 2.176% | 2.608% | 1.083% | 1.944% | 2.915%
| PDM3 |  | 1.351% | -7.902% | 1.526% | -1.313% | 9.504% | 0.2722% | -4.293% | -0.3088% | 0.7288% | -18.53% | -0.4748% | 9.527% | 10.74% | -6.783% | 2.381% | -3.397% | -19.45% | 2.098% | -10.27% | -5.563% | 4.041% | 16.65% | -31.02%
| Z-Score (Absolute) |  | 0.26 | 1.18 | 0.28 | 0.16 | 1.52 | 0.09 | 0.62 | 0 | 0.16 | 2.82 | 0.03 | 1.52 | 1.71 | 1 | 0.42 | 0.48 | 2.97 | 0.37 | 1.54 | 0.81 | 0.67 | 2.63 | 4.76
| Analytical results for Tm in OREAS 313 (Certified Value 0.18 ppm) | [26 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.12 | 2.16 | 2.17 | 2.20 | [18 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [18 empty] | 2
2
| 1 | 1 | 0.1431 | 0.196 | 0.21179 | 0.3 | 0.2 | 0.2 | 0.14 | [18 empty] | 1
| 1 | 2 | 0.15 | 0.201 | 0.22535 | 0.3 | 0.19 | 0.2 | 0.14 | [18 empty] | 8
| 1 | 3 | 0.1419 | 0.198 | 0.2232 | 0.3 | 0.21 | 0.2 | 0.14 | [18 empty] | 16
| 1 | 4 | 0.1356 | 0.203 | 0.22322 | 0.3 | 0.2 | 0.2 | 0.14 | [18 empty] | 0.17925
| 1 | 5 | 0.1329 | 0.195 | 0.20913 | 0.3 | 0.2 | 0.2 | 0.15 | [18 empty] | 57
| 1 | 6 | 0.1327 | 0.196 | 0.21953 | 0.3 | 0.19 | 0.2 | 0.14
| Mean |  | 0.13937 | 0.19817 | 0.2187 | 0.3 | 0.19833 | 0.2 | 0.14167
| Median |  | 0.13875 | 0.197 | 0.22137 | 0.3 | 0.2 | 0.2 | 0.14
| Std Dev. |  | 0.0068392 | 0.0031885 | 0.0067057 | 0 | 0.0075277 | 3.0405e-17 | 0.0040825
| Rel.Std.Dev. |  | 4.907% | 1.609% | 3.066% | 0% | 3.795% | 1.52e-14% | 2.882%
| PDM3 |  | -22.25% | 10.55% | 22.01% | 67.37% | 10.65% | 11.58% | -20.97%
| Z-Score (Absolute) |  | 1.93 | 0 | 0.67 | NA | 0.01 | NA | 1.86
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 313 (Certified Value 3.45 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.24 | 2.9084 | 3.5 | < 5 | 3.61 | 3.4733 | 3.6 | 3 | 3.5 | 3.5 | 3.2 | 3.7 | 3.44 | 4.0702 | 3.14 | 3.3 | 3.6 | 3.27 | 3.6 | 3 | 3.15 | 3.48 | 3.71 | 3.7 | 3.2 | 1
| 1 | 2 | 3.32 | 2.9602 | 3.6 | < 5 | 3.52 | 3.6053 | 3.6 | 3.1 | 3.3 | 3.9 | 3.2 | 3.7 | 3.59 | 3.9415 | 3.28 | 3.27 | 3.7 | 3.32 | 3.5 | 2.9 | 3.17 | 3.51 | 3.75 | 3.6 | 3.3 | 17
| 1 | 3 | 3.32 | 2.9329 | 3.4 | < 5 | 3.66 | 3.5384 | 3.5 | 3.1 | 3.3 | 3.6 | 3.3 | 3.6 | 3.65 | 3.9421 | 3.22 | 3.36 | 3.5 | 3.1 | 3.6 | 2.9 | 3.11 | 3.8 | 3.71 | 3.6 | 3.2 | 16
| 1 | 4 | 3.33 | 2.8212 | 3.6 | < 5 | 3.61 | 3.6653 | 3.6 | 3 | 3.5 | 3.6 | 3.2 | 3.6 | 3.4 | 4.0804 | 3.11 | 3.33 | 3.5 | 3.15 | 3.6 | 2.8 | 3.1 | 3.54 | 3.73 | 3.6 | 3.3 | 3.4476
| 1 | 5 | 3.3 | 2.855 | 3.4 | < 5 | 3.59 | 3.5519 | 3.8 | 3.3 | 3.5 | 3.5 | 3.3 | 3.8 | 3.52 | 3.8988 | 3.23 | 3.41 | 3.6 | 3.28 | 3.3 | 2.9 | 3.3 | 3.46 | 3.69 | 3.6 | 3.1 | 58
| 1 | 6 | 3.31 | 2.9158 | 3.6 | < 5 | 3.65 | 3.4662 | 3.7 | 3.1 | 3.4 | 3.5 | 3.3 | 3.7 | 3.63 | 4.0063 | 3.2 | 3.36 | 3.6 | 3.16 | 3.7 | 2.8 | 3.22 | 3.86 | 3.72 | 3.6 | 3.2
| Mean |  | 3.3033 | 2.8989 | 3.5167 |  | 3.6067 | 3.5501 | 3.6333 | 3.1 | 3.4167 | 3.6 | 3.25 | 3.6833 | 3.5383 | 3.9899 | 3.1967 | 3.3383 | 3.5833 | 3.2133 | 3.55 | 2.8833 | 3.175 | 3.6083 | 3.7183 | 3.6167 | 3.2167
| Median |  | 3.315 | 2.9121 | 3.55 |  | 3.61 | 3.5452 | 3.6 | 3.1 | 3.45 | 3.55 | 3.25 | 3.7 | 3.555 | 3.9742 | 3.21 | 3.345 | 3.6 | 3.215 | 3.6 | 2.9 | 3.16 | 3.525 | 3.715 | 3.6 | 3.2
| Std Dev. |  | 0.03266 | 0.051486 | 0.098319 |  | 0.050067 | 0.076668 | 0.10328 | 0.10954 | 0.098319 | 0.15492 | 0.054772 | 0.075277 | 0.10265 | 0.074611 | 0.062183 | 0.049565 | 0.075277 | 0.088015 | 0.13784 | 0.075277 | 0.075033 | 0.17486 | 0.020412 | 0.040825 | 0.075277
| Rel.Std.Dev. |  | 0.9887% | 1.776% | 2.796% |  | 1.388% | 2.16% | 2.843% | 3.534% | 2.878% | 4.303% | 1.685% | 2.044% | 2.901% | 1.87% | 1.945% | 1.485% | 2.101% | 2.739% | 3.883% | 2.611% | 2.363% | 4.846% | 0.549% | 1.129% | 2.34%
| PDM3 |  | -4.183% | -15.91% | 2.005% |  | 4.615% | 2.974% | 5.389% | -10.08% | -0.8961% | 4.422% | -5.73% | 6.839% | 2.633% | 15.73% | -7.277% | -3.168% | 3.938% | -6.794% | 2.971% | -16.37% | -7.906% | 4.663% | 7.854% | 4.905% | -6.697%
| Z-Score (Absolute) |  | 0.81 | 2.34 | 0 | 3.84 | 0.34 | 0.13 | 0.44 | 1.58 | 0.38 | 0.32 | 1.01 | 0.63 | 0.08 | 1.79 | 1.21 | 0.67 | 0.25 | 1.15 | 0.13 | 2.39 | 1.29 | 0.35 | 0.76 | 0.38 | 1.13
| Analytical results for V in OREAS 313 (Certified Value 133 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 133 | 132.8 | 146 | 138 | 128 | 132.2074 | 137 | 135 | 129 | 128 | 126 | 157 | 135 | 140.7978 | 136 | 137 | 130 | 138 | 135 | 130 | 129 | 129 | 129 | 136 | 129 | 1
| 1 | 2 | 132 | 130.28 | 141 | 140.7 | 125 | 134.886 | 137 | 137 | 128 | 130 | 128 | 160 | 135 | 136.0518 | 137 | 140 | 131 | 138 | 132 | 131 | 127 | 129 | 134 | 135 | 132 | 18
| 1 | 3 | 132 | 132.61 | 145 | 140.8 | 118 | 130.3482 | 138 | 141 | 130 | 131 | 128 | 145 | 141 | 136.9889 | 137 | 142 | 133 | 135 | 133 | 130 | 127 | 133 | 133 | 133 | 133 | 16
| 1 | 4 | 133 | 133.52 | 142 | 139.1 | 128 | 136.1069 | 135 | 136 | 129 | 133 | 128 | 137 | 135 | 139.6619 | 135 | 135 | 128 | 134 | 133 | 128 | 128 | 129 | 128 | 141 | 134 | 133.2304
| 1 | 5 | 130 | 131 | 141 | 136.7 | 123 | 129.5901 | 139 | 132 | 129 | 129 | 130 | 141 | 135 | 139.2047 | 134 | 142 | 132 | 137 | 124 | 131 | 128 | 130 | 132 | 134 | 127 | 59
| 1 | 6 | 130 | 130.28 | 140 | 139.1 | 129 | 132.8374 | 138 | 137 | 129 | 128 | 128 | 157 | 135 | 137.8023 | 134 | 139 | 132 | 134 | 132 | 130 | 129 | 130 | 134 | 133 | 130
| Mean |  | 131.6667 | 131.7483 | 142.5 | 139.0667 | 125.1667 | 132.6627 | 137.3333 | 136.3333 | 129 | 129.8333 | 128 | 149.5 | 136 | 138.4179 | 135.5 | 139.1667 | 131 | 136 | 131.5 | 130 | 128 | 130 | 131.6667 | 135.3333 | 130.8333
| Median |  | 132 | 131.805 | 141.5 | 139.1 | 126.5 | 132.5224 | 137.5 | 136.5 | 129 | 129.5 | 128 | 151 | 135 | 138.5035 | 135.5 | 139.5 | 131.5 | 136 | 132.5 | 130 | 128 | 129.5 | 132.5 | 134.5 | 131
| Std Dev. |  | 1.3663 | 1.4042 | 2.429 | 1.5756 | 4.1673 | 2.5235 | 1.3663 | 2.9439 | 0.63246 | 1.9408 | 1.2649 | 9.7108 | 2.4495 | 1.7804 | 1.3784 | 2.7869 | 1.7889 | 1.8974 | 3.8341 | 1.0954 | 0.89443 | 1.5492 | 2.582 | 3.0111 | 2.6394
| Rel.Std.Dev. |  | 1.038% | 1.066% | 1.705% | 1.133% | 3.329% | 1.902% | 0.9948% | 2.159% | 0.4903% | 1.495% | 0.9882% | 6.496% | 1.801% | 1.286% | 1.017% | 2.003% | 1.366% | 1.395% | 2.916% | 0.8427% | 0.6988% | 1.192% | 1.961% | 2.225% | 2.017%
| PDM3 |  | -1.174% | -1.112% | 6.958% | 4.381% | -6.052% | -0.4261% | 3.08% | 2.329% | -3.175% | -2.55% | -3.926% | 12.21% | 2.079% | 3.894% | 1.704% | 4.456% | -1.674% | 2.079% | -1.299% | -2.425% | -3.926% | -2.425% | -1.174% | 1.578% | -1.799%
| Z-Score (Absolute) |  | 0.01 | 0 | 1.93 | 1.32 | 1.18 | 0.16 | 1 | 0.82 | 0.49 | 0.34 | 0.67 | 3.19 | 0.76 | 1.2 | 0.67 | 1.33 | 0.13 | 0.76 | 0.04 | 0.31 | 0.67 | 0.31 | 0.01 | 0.64 | 0.16
| Analytical results for W in OREAS 313 (Certified Value 1.37 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 1.3 | 0.99711 | 2.3 | 1.74 | 1.4 | 1.4558 | 1.7 | 1.5 | 1.3 | 1.2 | 1.3 | 1.7 | 1.4 | 1.267 | 1.9 | 1.4 | 1.3 | 1.4 | 1.2 | 2 | 1.17 | 1.2 | 1.6 | 1.2 |  | 1
| 1 | 2 | 1.3 | 0.99651 | 2.4 | 1.84 | 1.3 | 1.5312 | 1.7 | 1.4 | 1.4 | 1.3 | 1.4 | 1.9 | 1.4 | 1.417 | 1.16 | 1.4 | 1.3 | 1.4 | 1.2 | 1 | 1.23 | 1.2 | 1.6 | 1.1 |  | 19
| 1 | 3 | 1.3 | 1.0425 | 2.3 | 1.86 | 1.4 | 1.5128 | 1.6 | 1.4 | 1.3 | 1.2 | 1.4 | 1.7 | 1.4 | 1.2551 | 1.69 | 1.4 | 1.3 | 1.4 | 1.1 | 2 | 1.2 | 1.3 | 1.5 | 1.2 |  | 16
| 1 | 4 | 1.3 | 0.98212 | 2.4 | 1.79 | 1.5 | 1.5213 | 1.6 | 1.4 | 1.4 | 1.2 | 1.3 | 1.5 | 1.3 | 1.1809 | 1.18 | 1.4 | 1.3 | 1.4 | 1.2 | 3 | 1.12 | 1.3 | 1.6 | 1.1 |  | 1.3698
| 1 | 5 | 1.3 | 1.0371 | 2.1 | 1.67 | 1.3 | 1.4616 | 1.6 | 1.3 | 1.2 | 1.3 | 1.4 | 1.8 | 1.4 | 1.3366 | 1.15 | 1.6 | 1.3 | 1.3 | 1.2 | 2 | 1.13 | 1.3 | 1.6 | 1 |  | 60
| 1 | 6 | 1.3 | 1.0588 | 2.4 | 1.48 | 1.3 | 1.4475 | 1.6 | 1.4 | 1.4 | 1.2 | 1.4 | 1.8 | 1.4 | 1.362 | 1.32 | 1.5 | 1.4 | 1.4 | 1.1 | 2 | 1.22 | 1.2 | 1.6 | 1.1
| Mean |  | 1.3 | 1.019 | 2.3167 | 1.73 | 1.3667 | 1.4883 | 1.6333 | 1.4 | 1.3333 | 1.2333 | 1.3667 | 1.7333 | 1.3833 | 1.3031 | 1.4 | 1.45 | 1.3167 | 1.3833 | 1.1667 | 2 | 1.1783 | 1.25 | 1.5833 | 1.1167
| Median |  | 1.3 | 1.0171 | 2.35 | 1.765 | 1.35 | 1.4872 | 1.6 | 1.4 | 1.35 | 1.2 | 1.4 | 1.75 | 1.4 | 1.3018 | 1.25 | 1.4 | 1.3 | 1.4 | 1.2 | 2 | 1.185 | 1.25 | 1.6 | 1.1
| Std Dev. |  | 0 | 0.031004 | 0.1169 | 0.14057 | 0.08165 | 0.037338 | 0.05164 | 0.063246 | 0.08165 | 0.05164 | 0.05164 | 0.13663 | 0.040825 | 0.084976 | 0.31906 | 0.083666 | 0.040825 | 0.040825 | 0.05164 | 0.63246 | 0.046224 | 0.054772 | 0.040825 | 0.075277
| Rel.Std.Dev. |  | 0% | 3.043% | 5.046% | 8.125% | 5.974% | 2.509% | 3.162% | 4.518% | 6.124% | 4.187% | 3.779% | 7.882% | 2.951% | 6.521% | 22.79% | 5.77% | 3.101% | 2.951% | 4.426% | 31.62% | 3.923% | 4.382% | 2.578% | 6.741%
| PDM3 |  | -5.095% | -25.61% | 69.13% | 26.3% | -0.2283% | 8.655% | 19.24% | 2.205% | -2.662% | -9.962% | -0.2283% | 26.54% | 0.9884% | -4.869% | 2.205% | 5.855% | -3.878% | 0.9884% | -14.83% | 46.01% | -13.98% | -8.745% | 15.59% | -18.48%
| Z-Score (Absolute) |  | 0.42 | 2.01 | 5.33 | 2.01 | 0.05 | 0.64 | 1.46 | 0.14 | 0.24 | 0.8 | 0.05 | 2.03 | 0.05 | 0.41 | 0.14 | 0.42 | 0.33 | 0.05 | 1.18 | 3.54 | 1.11 | 0.71 | 1.18 | 1.46
| Analytical results for Y in OREAS 313 (Certified Value 13.5 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
1
| 1 | 1 | 12.5 | 12.58 | 10.9 | 14.8 | 14.33 | 14.6213 | 13.7 | 13.4 | 13.6 | 14 | 12.4 | 20.4 | 10 | 18.0187 | 10.8 | 12.13 | 13.2 | 14.8 | 12 | 14.1 | 14.6 | 14.28 | 17.2 | 13.2 |  | 1
| 1 | 2 | 12.7 | 12.5676 | 11.1 | 15.7 | 12.2 | 14.7743 | 14.5 | 13.6 | 14.3 | 14.2 | 12.5 | 19.7 | 10.2 | 17.6781 | 11.6 | 12.5 | 13.4 | 15.1 | 12 | 14.6 | 15 | 14.65 | 17.2 | 13.7 |  | 20
| 1 | 3 | 12.6 | 12.6765 | 11.1 | 15.6 | 12.04 | 14.3535 | 12.9 | 13.6 | 14.1 | 14.4 | 13.2 | 19.1 | 10.9 | 17.6717 | 12.8 | 12.41 | 13 | 15.2 | 12 | 13 | 14.9 | 15.8 | 17.2 | 15.4 |  | 16
| 1 | 4 | 13.3 | 12.6791 | 10.4 | 15.6 | 12.89 | 15.0917 | 14.8 | 13.4 | 13.7 | 14.3 | 12.7 | 18.1 | 10.7 | 18.2614 | 10.8 | 12.26 | 12.9 | 15.6 | 12 | 12.9 | 14.1 | 14.11 | 17.1 | 13.9 |  | 13.499
| 1 | 5 | 12.7 | 12.5773 | 11.4 | 14.8 | 12.21 | 14.1567 | 14.8 | 12.6 | 15 | 13.8 | 13 | 19 | 10.2 | 18.2852 | 11.5 | 12.49 | 13 | 14.8 | 12 | 14.3 | 14.5 | 14.6 | 16.8 | 13.1 |  | 61
| 1 | 6 | 12.2 | 12.7124 | 10.7 | 14.2 | 12.75 | 14.6098 | 14.8 | 13.2 | 15.1 | 14 | 13 | 18.6 | 10.4 | 18.1553 | 11 | 12.43 | 13.4 | 15 | 12 | 13.2 | 15.8 | 14.48 | 16.8 | 13.8
| Mean |  | 12.6667 | 12.6322 | 10.9333 | 15.1167 | 12.7367 | 14.6012 | 14.25 | 13.3 | 14.3 | 14.1167 | 12.8 | 19.15 | 10.4 | 18.0117 | 11.4167 | 12.37 | 13.15 | 15.0833 | 12 | 13.6833 | 14.8167 | 14.6533 | 17.05 | 13.85
| Median |  | 12.65 | 12.6283 | 11 | 15.2 | 12.48 | 14.6156 | 14.65 | 13.4 | 14.2 | 14.1 | 12.85 | 19.05 | 10.3 | 18.087 | 11.25 | 12.42 | 13.1 | 15.05 | 12 | 13.65 | 14.75 | 14.54 | 17.15 | 13.75
| Std Dev. |  | 0.36148 | 0.064029 | 0.35024 | 0.608 | 0.85007 | 0.32557 | 0.78677 | 0.37417 | 0.63561 | 0.22286 | 0.31623 | 0.81179 | 0.34059 | 0.27738 | 0.76004 | 0.14574 | 0.21679 | 0.29944 | 0 | 0.73598 | 0.57764 | 0.59705 | 0.19748 | 0.82644
| Rel.Std.Dev. |  | 2.854% | 0.5069% | 3.203% | 4.022% | 6.674% | 2.23% | 5.521% | 2.813% | 4.445% | 1.579% | 2.471% | 4.239% | 3.275% | 1.54% | 6.657% | 1.178% | 1.649% | 1.985% | 0% | 5.379% | 3.899% | 4.074% | 1.158% | 5.967%
| PDM3 |  | -6.166% | -6.421% | -19.01% | 11.98% | -5.647% | 8.165% | 5.563% | -1.474% | 5.934% | 4.576% | -5.178% | 41.86% | -22.96% | 33.43% | -15.43% | -8.363% | -2.585% | 11.74% | -11.1% | 1.366% | 9.761% | 8.551% | 26.31% | 2.6%
| Z-Score (Absolute) |  | 0.72 | 0.74 | 1.77 | 0.77 | 0.67 | 0.45 | 0.24 | 0.33 | 0.27 | 0.16 | 0.64 | 3.21 | 2.09 | 2.52 | 1.47 | 0.9 | 0.42 | 0.75 | NA | 0.1 | 0.59 | 0.49 | 1.94 | 0
Results from laboratory 2.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 313 (Certified Value 1.45 ppm) | [26 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.12 | 2.13 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 | [14 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [14 empty] | 2
2
| 1 | 1 | 1.2067 | 1.44 | 1.81 | 1.4652 | 2 | 1.1 | 1.42 | 1.3 | 0.99 | 1.7 | 1.45 | [14 empty] | 1
| 1 | 2 | 1.2493 | 1.42 | 1.96 | 1.518 | 2 | 1.1 | 1.38 | 1.3 | 0.99 | 1.7 | 1.53 | [14 empty] | 21
| 1 | 3 | 1.2821 | 1.444 | 1.97 | 1.4307 | 2 | 1.2 | 1.49 | 1.3 | 1.04 | 1.6 | 1.51 | [14 empty] | 16
| 1 | 4 | 1.2471 | 1.42 | 1.92 | 1.5622 | 1.9 | 1.2 | 1.29 | 1.2 | 1.01 | 1.6 | 1.46 | [14 empty] | 1.4549
| 1 | 5 | 1.2682 | 1.42 | 1.74 | 1.4555 | 2 | 1.2 | 1.42 | 1.2 | 1.06 | 1.7 | 1.51 | [14 empty] | 62
| 1 | 6 | 1.2772 | 1.42 | 1.79 | 1.4939 | 1.9 | 1.2 | 1.33 | 1.3 | 1.02 | 1.6 | 1.61
| Mean |  | 1.2551 | 1.4273 | 1.865 | 1.4876 | 1.9667 | 1.1667 | 1.3883 | 1.2667 | 1.0183 | 1.65 | 1.5117
| Median |  | 1.2588 | 1.42 | 1.865 | 1.4796 | 2 | 1.2 | 1.4 | 1.3 | 1.015 | 1.65 | 1.51
| Std Dev. |  | 0.027674 | 0.011431 | 0.097314 | 0.047561 | 0.05164 | 0.05164 | 0.071391 | 0.05164 | 0.027869 | 0.054772 | 0.057417
| Rel.Std.Dev. |  | 2.205% | 0.8009% | 5.218% | 3.197% | 2.626% | 4.426% | 5.142% | 4.077% | 2.737% | 3.32% | 3.798%
| PDM3 |  | -13.73% | -1.891% | 28.19% | 2.25% | 35.18% | -19.81% | -4.572% | -12.93% | -30% | 13.41% | 3.905%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.71 | 0.24 | 2.11 | 1.02 | 0.15 | 0.63 | 1.6 | 0.87 | 0.33
| Analytical results for Zn in OREAS 313 (Certified Value 3.55 wt.%) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*MS | 4A*OES | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 3.65 | 3.4776 | 3.4496 | 3.79 | 3.5121 | 3.5622 | 3.45 | 3.51 | 3.56 | 3.48 | 3.44 | 3.57 | 3.56 | 3.619 | 3.03 | 3.5415 | - | 3.6149 | 3.58 | 3.62 | 3.66 | 3.47 | 3.6636 | 3.46 | 3.43 | 1
| 1 | 2 | 3.64 | 3.4956 | 3.42 | 3.73 | 3.5769 | 3.5223 | 3.43 | 3.54 | 3.54 | 3.5 | 3.45 | 3.62 | 3.44 | 3.5812 | 2.93 | 3.4653 | - | 3.7004 | 3.61 | 3.54 | 3.61 | 3.54 | 3.5687 | 3.47 | 3.43 | #N/A
| 1 | 3 | 3.64 | 3.4649 | 3.46 | 3.71 | 3.5113 | 3.5778 | 3.44 | 3.55 | 3.55 | 3.5 | 3.45 | 3.75 | 3.45 | 3.5957 | 3.01 | 3.5564 | - | 3.6255 | 3.63 | 3.52 | 3.62 | 3.51 | 3.6027 | 3.45 | 3.42 | 16
| 1 | 4 | 3.58 | 3.4872 | 3.465 | 3.72 | 3.4876 | 3.599 | 3.45 | 3.53 | 3.53 | 3.51 | 3.44 | 3.64 | 3.4 | 3.527 | 2.84 | 3.5713 | - | 3.6174 | 3.63 | 3.65 | 3.69 | 3.49 | 3.7017 | 3.45 | 3.42 | 3.5482
| 1 | 5 | 3.58 | 3.4854 | 3.46 | 3.75 | 3.5494 | 3.5496 | 3.52 | 3.51 | 3.53 | 3.49 | 3.4 | 3.65 | 3.26 | 3.5797 | 3.2 | 3.6294 | - | 3.6168 | 3.63 | 3.56 | 3.66 | 3.44 | 3.6664 | 3.5 | 3.46 | 63
| 1 | 6 | 3.63 | 3.4768 | 3.46 | 3.73 | 3.553 | 3.5411 | 3.5 | 3.5 | 3.45 | 3.51 | 3.44 | 3.75 | 3.42 | 3.5134 | 3 | 3.6013 | - | 3.6453 | 3.59 | 3.59 | 3.69 | 3.49 | 3.7336 | 3.47 | 3.43
| Mean |  | 3.62 | 3.4813 | 3.4524 | 3.7383 | 3.5317 | 3.5587 | 3.465 | 3.5233 | 3.5267 | 3.4983 | 3.4367 | 3.6633 | 3.4217 | 3.5693 | 3.0017 | 3.5609 |  | 3.6367 | 3.6117 | 3.58 | 3.655 | 3.49 | 3.6561 | 3.4667 | 3.4317
| Median |  | 3.635 | 3.4815 | 3.46 | 3.73 | 3.5307 | 3.5559 | 3.45 | 3.52 | 3.535 | 3.5 | 3.44 | 3.645 | 3.43 | 3.5804 | 3.005 | 3.5639 |  | 3.6215 | 3.62 | 3.575 | 3.66 | 3.49 | 3.665 | 3.465 | 3.43
| Std Dev. |  | 0.031623 | 0.01059 | 0.016671 | 0.028577 | 0.033332 | 0.027271 | 0.036194 | 0.019664 | 0.039328 | 0.01169 | 0.018619 | 0.072572 | 0.096833 | 0.040833 | 0.11957 | 0.056513 |  | 0.033169 | 0.022286 | 0.049396 | 0.033912 | 0.034059 | 0.061218 | 0.018619 | 0.01472
| Rel.Std.Dev. |  | 0.8736% | 0.3042% | 0.4829% | 0.7644% | 0.9438% | 0.7663% | 1.045% | 0.5581% | 1.115% | 0.3342% | 0.5418% | 1.981% | 2.83% | 1.144% | 3.983% | 1.587% |  | 0.9121% | 0.6171% | 1.38% | 0.9278% | 0.9759% | 1.674% | 0.5371% | 0.4289%
| PDM3 |  | 2.025% | -1.886% | -2.698% | 5.36% | -0.4636% | 0.2962% | -2.344% | -0.6998% | -0.6059% | -1.404% | -3.142% | 3.246% | -3.565% | 0.5966% | -15.4% | 0.358% |  | 2.496% | 1.79% | 0.8972% | 3.011% | -1.639% | 3.042% | -2.297% | -3.283%
| Z-Score (Absolute) |  | 0.87 | 0.46 | 0.73 | 2 | 0.02 | 0.28 | 0.61 | 0.06 | 0.02 | 0.3 | 0.89 | 1.28 | 1.03 | 0.38 | 5.05 | 0.3 | NA | 1.03 | 0.79 | 0.49 | 1.2 | 0.37 | 1.21 | 0.6 | 0.93
| Analytical results for Zr in OREAS 313 (Certified Value 70 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 68.9 | 66.93 | 67.4 | 59.27 | 68.6 | 73.8092 | 76.1 | 71.3 | 69.3 | 68 | 63.8 | 78 | 66.4 | 74.8336 | 67.8 | 71.3 | 74.7 | 70.8 | 66 | 72.5 | 65.1 | 71.3 | 74.8 | 71.3 |  | 1
| 1 | 2 | 67.6 | 64.96 | 66.6 | 59.16 | 64.7 | 74.2455 | 74.4 | 72.7 | 69 | 72.8 | 66.9 | 80 | 66.8 | 73.9359 | 67.4 | 74.3 | 74.2 | 69.6 | 67 | 73.9 | 64.4 | 73 | 76.5 | 72.1 |  | 22
| 1 | 3 | 68 | 68.26 | 62.4 | 61.71 | 66.1 | 71.893 | 74.7 | 73.2 | 68.6 | 70.5 | 67.1 | 75 | 69.1 | 74.8391 | 68.3 | 69.2 | 74.4 | 70.2 | 65 | 73.5 | 66.8 | 71.5 | 74.8 | 70.9 |  | 16
| 1 | 4 | 68.6 | 63.65 | 68.2 | 61.4 | 68.1 | 74.3577 | 72.2 | 72.4 | 68.1 | 71.2 | 66.1 | 72 | 67.9 | 75.4733 | 65.2 | 69.2 | 72.1 | 72 | 64 | 72.1 | 64.1 | 70.9 | 75.7 | 71.2 |  | 69.8741
| 1 | 5 | 67.4 | 68.91 | 64.4 | 60.96 | 64.6 | 70.8597 | 74.1 | 68.1 | 69.4 | 67.2 | 68.1 | 75 | 66.7 | 72.9789 | 69.4 | 72 | 73.1 | 66.6 | 67 | 73.7 | 64.3 | 74 | 75.6 | 71.1 |  | 64
| 1 | 6 | 67.7 | 66.03 | 65.6 | 60.38 | 67 | 73.2865 | 76.6 | 72.7 | 70.6 | 65.8 | 67.5 | 78 | 67.4 | 73.2438 | 67.5 | 71.6 | 74.1 | 72.2 | 66 | 73 | 67.5 | 72.7 | 74.4 | 71.2
| Mean |  | 68.0333 | 66.4567 | 65.7667 | 60.48 | 66.5167 | 73.0753 | 74.6833 | 71.7333 | 69.1667 | 69.25 | 66.5833 | 76.3333 | 67.3833 | 74.2174 | 67.6 | 71.2667 | 73.7667 | 70.2333 | 65.8333 | 73.1167 | 65.3667 | 72.2333 | 75.3 | 71.3
| Median |  | 67.85 | 66.48 | 66.1 | 60.67 | 66.55 | 73.5478 | 74.55 | 72.55 | 69.15 | 69.25 | 67 | 76.5 | 67.1 | 74.3847 | 67.65 | 71.45 | 74.15 | 70.5 | 66 | 73.25 | 64.75 | 72.1 | 75.2 | 71.2
| Std Dev. |  | 0.59554 | 1.9897 | 2.1219 | 1.0776 | 1.6869 | 1.4074 | 1.5664 | 1.8896 | 0.8501 | 2.6696 | 1.5158 | 2.8752 | 0.99883 | 0.99041 | 1.3842 | 1.918 | 0.97912 | 2.0452 | 1.169 | 0.7111 | 1.439 | 1.1928 | 0.7746 | 0.41473
| Rel.Std.Dev. |  | 0.8754% | 2.994% | 3.226% | 1.782% | 2.536% | 1.926% | 2.097% | 2.634% | 1.229% | 3.855% | 2.277% | 3.767% | 1.482% | 1.334% | 2.048% | 2.691% | 1.327% | 2.912% | 1.776% | 0.9726% | 2.201% | 1.651% | 1.029% | 0.5817%
| PDM3 |  | -2.634% | -4.891% | -5.878% | -13.44% | -4.805% | 4.581% | 6.883% | 2.661% | -1.012% | -0.8932% | -4.71% | 9.244% | -3.565% | 6.216% | -3.255% | 1.993% | 5.571% | 0.5141% | -5.783% | 4.641% | -6.451% | 3.376% | 7.765% | 2.041%
| Z-Score (Absolute) |  | 0.35 | 0.68 | 0.82 | 1.92 | 0.67 | 0.69 | 1.02 | 0.41 | 0.12 | 0.1 | 0.65 | 1.37 | 0.49 | 0.93 | 0.44 | 0.32 | 0.83 | 0.1 | 0.81 | 0.7 | 0.91 | 0.52 | 1.15 | 0.32
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C930: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1095: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 313 (Certified Value 74.6 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 73.1 | 75.884 | 79.58 | 73.8059 | 75.8 | 77.6 | 80.3 | 75.9 | 73.8 | 68.8 | 77.42 | 74.3748 | 73.6 | 72.774 | 73.8 | 74 | 73 | 74.8 | 76.1 | 70.47 | 73.43 | 61.2 |  |  | 1
| 1 | 2 | 71 | 73.4995 | 78.89 | 74.441 | 76.6 | 79.1 | 77.7 | 75 | 75.4 | 71.2 | 75.93 | 75.3854 | 74.4 | 73.26 | 73.4 | 74 | 72 | 74.7 | 75.6 | 70.26 | 73.42 | 60.4 |  |  | 10
| 1 | 3 | 73 | 74.3325 | 79.11 | 74.3756 | 76.1 | 77.5 | 78.2 | 76.1 | 74.4 | 70.4 | 74.94 | 73.2425 | 70.7 | 74.286 | 74.5 | 75 | 72 | 74.2 | 74.5 | 69.95 | 74.19 | 60.1 |  |  | 16
| 1 | 4 | 70.6 | 76.959 | 78.35 | 74.7808 | 76.1 | 78.6 | 80.2 | 75.5 | 74.2 | 70 | 77.19 | 72.9836 | 73.4 | 75.438 | 73.9 | 74 | 72 | 72.7 | 74.9 | 69.37 | 74.65 | 59.7 |  |  | 74.5965
| 1 | 5 | 75.7 | 76.2535 | 79.09 | 74.7361 | 75 | 77.4 | 79.8 | 76.2 | 75.2 | 70 | 76.24 | 74.6393 | 78.8 | 72.396 | 75.2 | 74 | 72 | 74.1 | 74.6 | 67.27 | 73.48 | 61 |  |  | 66
| 1 | 6 | 72.6 | 75.8785 | 79.22 | 74.6593 | 75.8 | 78 | 79.1 | 75.9 | 74.5 | 69.3 | 77.38 | 73.0524 | 75.3 | 71.622 | 72.6 | 74 | 71 | 74 | 74.9 | 70.78 | 74.19 | 60.3
| Mean |  | 72.6667 | 75.4678 | 79.04 | 74.4665 | 75.9 | 78.0333 | 79.2167 | 75.7667 | 74.5833 | 69.95 | 76.5167 | 73.9463 | 74.3667 | 73.296 | 73.9 | 74.1667 | 72 | 74.0833 | 75.1 | 69.6833 | 73.8933 | 60.45
| Median |  | 72.8 | 75.8812 | 79.1 | 74.5502 | 75.95 | 77.8 | 79.45 | 75.9 | 74.45 | 70 | 76.715 | 73.8087 | 74 | 73.017 | 73.85 | 74 | 72 | 74.15 | 74.9 | 70.105 | 73.835 | 60.35
| Std Dev. |  | 1.8195 | 1.292 | 0.40743 | 0.36165 | 0.52915 | 0.68313 | 1.0797 | 0.44572 | 0.608 | 0.83845 | 0.9921 | 0.9956 | 2.6643 | 1.3751 | 0.89443 | 0.40825 | 0.63246 | 0.75211 | 0.6229 | 1.2765 | 0.52118 | 0.56125
| Rel.Std.Dev. |  | 2.504% | 1.712% | 0.5155% | 0.4857% | 0.6972% | 0.8754% | 1.363% | 0.5883% | 0.8152% | 1.199% | 1.297% | 1.346% | 3.583% | 1.876% | 1.21% | 0.5504% | 0.8784% | 1.015% | 0.8294% | 1.832% | 0.7053% | 0.9285%
| PDM3 |  | -2.587% | 1.168% | 5.957% | -0.1743% | 1.747% | 4.607% | 6.194% | 1.569% | -0.01762% | -6.229% | 2.574% | -0.8715% | -0.3081% | -1.743% | -0.9337% | -0.5762% | -3.481% | -0.6879% | 0.675% | -6.586% | -0.9426% | -18.96%
| Z-Score (Absolute) |  | 0.8 | 0.6 | 2.38 | 0.1 | 0.82 | 1.88 | 2.47 | 0.75 | 0.16 | 2.16 | 1.12 | 0.16 | 0.05 | 0.48 | 0.18 | 0.05 | 1.13 | 0.09 | 0.42 | 2.29 | 0.19 | 6.9
| Analytical results for Al in OREAS 313 (Certified Value 0.787 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 1.05 | 0.68002 | 0.75 | 0.724 | 0.86046 | 0.79 | 0.77 | 0.83 | 0.75 | 0.77 | 0.82 | 0.793 | 0.87371 | 0.75 | 0.79 | 0.81 | 0.79 | 0.75 | 0.86 | 0.83 | 0.69 |  |  |  | 1
| 1 | 2 | 1.05 | 0.68637 | 0.71 | 0.754 | 0.85784 | 0.8 | 0.78 | 0.82 | 0.74 | 0.78 | 0.82 | 0.75 | 0.85988 | 0.77 | 0.81 | 0.82 | 0.77 | 0.75 | 0.86 | 0.84 | 0.7 |  |  |  | #N/A
| 1 | 3 | 1.05 | 0.68755 | 0.7 | 0.75 | 0.86167 | 0.78 | 0.77 | 0.82 | 0.74 | 0.78 | 0.85 | 0.759 | 0.93221 | 0.76 | 0.78 | 0.71 | 0.78 | 0.75 | 0.86 | 0.79 | 0.69 |  |  |  | 16
| 1 | 4 | 1.05 | 0.69975 | 0.79 | 0.735 | 0.83585 | 0.81 | 0.78 | 0.85 | 0.74 | 0.78 | 0.84 | 0.789 | 0.92552 | 0.78 | 0.87 | 0.86 | 0.79 | 0.75 | 0.85 | 0.82 | 0.68 |  |  |  | 0.78732
| 1 | 5 | 1.06 | 0.69738 | 0.74 | 0.758 | 0.86813 | 0.78 | 0.78 | 0.84 | 0.74 | 0.79 | 0.9 | 0.783 | 0.91595 | 0.84 | 0.73 | 0.79 | 0.77 | 0.76 | 0.85 | 0.83 | 0.68 |  |  |  | 67
| 1 | 6 | 1.03 | 0.68141 | 0.74 | 0.765 | 0.88379 | 0.79 | 0.78 | 0.84 | 0.75 | 0.78 | 0.84 | 0.748 | 0.88043 | 0.79 | 0.71 | 0.81 | 0.76 | 0.75 | 0.85 | 0.86 | 0.7
| Mean |  | 1.0483 | 0.68875 | 0.73833 | 0.74767 | 0.86129 | 0.79167 | 0.77667 | 0.83333 | 0.74333 | 0.78 | 0.845 | 0.77033 | 0.89795 | 0.78167 | 0.78167 | 0.8 | 0.77667 | 0.75167 | 0.855 | 0.82833 | 0.69
| Median |  | 1.05 | 0.68696 | 0.74 | 0.752 | 0.86107 | 0.79 | 0.78 | 0.835 | 0.74 | 0.78 | 0.84 | 0.771 | 0.89819 | 0.775 | 0.785 | 0.81 | 0.775 | 0.75 | 0.855 | 0.83 | 0.69
| Std Dev. |  | 0.0098319 | 0.0081548 | 0.031885 | 0.015319 | 0.015578 | 0.01169 | 0.005164 | 0.012111 | 0.005164 | 0.0063246 | 0.029496 | 0.020314 | 0.030335 | 0.031885 | 0.057417 | 0.0498 | 0.012111 | 0.0040825 | 0.0054772 | 0.023166 | 0.0089443
| Rel.Std.Dev. |  | 0.9379% | 1.184% | 4.319% | 2.049% | 1.809% | 1.477% | 0.6649% | 1.453% | 0.6947% | 0.8108% | 3.491% | 2.637% | 3.378% | 4.079% | 7.345% | 6.225% | 1.559% | 0.5431% | 0.6406% | 2.797% | 1.296%
| PDM3 |  | 33.15% | -12.52% | -6.221% | -5.036% | 9.396% | 0.5526% | -1.353% | 5.845% | -5.586% | -0.9292% | 7.327% | -2.157% | 14.05% | -0.7176% | -0.7176% | 1.611% | -1.353% | -4.528% | 8.597% | 5.21% | -12.36%
| Z-Score (Absolute) |  | 4.69 | 1.63 | 0.76 | 0.6 | 1.4 | 0.18 | 0.09 | 0.91 | 0.67 | 0.03 | 1.11 | 0.2 | 2.05 | 0 | 0 | 0.32 | 0.09 | 0.53 | 1.29 | 0.82 | 1.61
| Analytical results for As in OREAS 313 (Certified Value 93 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 1
| 1 | 1 | 85 | 98.37 | 109.8 | 82.9 | 89.8099 | 97.6 | 94.2 | 96.3 | 92.2 | 97.9 | 88 | 97 | 90.4408 | 91 | 89.9 | 90.2 | 103 | 87 | 102 | 88.3 | 87 | 98 | 77.5 |  | 1
| 1 | 2 | 86 | 97.16 | 109.41 | 83.1 | 87.0444 | 98.2 | 95.6 | 97 | 91.8 | 99.2 | 86.5 | 99 | 88.0181 | 92 | 92.3 | 88.8 | 105.1 | 86 | 101 | 97.1 | 90 | 99.5 | 76 |  | 11
| 1 | 3 | 86 | 97.37 | 112.6 | 77.8 | 89.0383 | 99.4 | 92.4 | 95.6 | 91.8 | 98.9 | 88.2 | 98 | 94.3025 | 101 | 91.4 | 91.4 | 88.8 | 87 | 102 | 88.2 | 87 | 94.3 | 76.1 |  | 16
| 1 | 4 | 85 | 98.34 | 109.2 | 90 | 90.5891 | 96.5 | 92.4 | 104.5 | 90.9 | 99.1 | 86.1 | 103 | 92.7044 | 96 | 94 | 90 | 109.8 | 86 | 102 | 88.4 | 86 | 99.7 | 76.1 |  | 93.3309
| 1 | 5 | 86 | 98.5 | 110.87 | 85.2 | 90.0609 | 97.5 | 93.6 | 104.5 | 89.9 | 98.3 | 90.7 | 94 | 89.2665 | 96 | 89.6 | 91 | 100.6 | 87 | 100 | 89.7 | 84 | 98 | 77.3 |  | 68
| 1 | 6 | 86 | 93.64 | 110 | 86.5 | 92.5358 | 97.7 | 92.5 | 97.9 | 90.8 | 99 | 88.4 | 92 | 93.5069 | 98 | 87.2 | 89.6 | 102.4 | 86 | 101 | 94.8 | 87 | 100.4 | 77.3
| Mean |  | 85.6667 | 97.23 | 110.3133 | 84.25 | 89.8464 | 97.8167 | 93.45 | 99.3 | 91.2333 | 98.7333 | 87.9833 | 97.1667 | 91.3732 | 95.6667 | 90.7333 | 90.1667 | 101.6167 | 86.5 | 101.3333 | 91.0833 | 86.8333 | 98.3167 | 76.7167
| Median |  | 86 | 97.855 | 109.9 | 84.15 | 89.9354 | 97.65 | 93.05 | 97.45 | 91.35 | 98.95 | 88.1 | 97.5 | 91.5726 | 96 | 90.65 | 90.1 | 102.7 | 86.5 | 101.5 | 89.05 | 87 | 98.75 | 76.7
| Std Dev. |  | 0.5164 | 1.847 | 1.2614 | 4.0933 | 1.8076 | 0.95376 | 1.2896 | 4.0993 | 0.85479 | 0.5164 | 1.6339 | 3.8687 | 2.5088 | 3.7238 | 2.3678 | 0.94163 | 7.0275 | 0.54772 | 0.8165 | 3.8778 | 1.9408 | 2.1904 | 0.71671
| Rel.Std.Dev. |  | 0.6028% | 1.9% | 1.143% | 4.859% | 2.012% | 0.9751% | 1.38% | 4.128% | 0.9369% | 0.523% | 1.857% | 3.981% | 2.746% | 3.892% | 2.61% | 1.044% | 6.916% | 0.6332% | 0.8058% | 4.257% | 2.235% | 2.228% | 0.9342%
| PDM3 |  | -8.212% | 4.178% | 18.2% | -9.73% | -3.734% | 4.806% | 0.1276% | 6.396% | -2.247% | 5.788% | -5.73% | 4.11% | -2.098% | 2.503% | -2.783% | -3.39% | 8.878% | -7.319% | 8.574% | -2.408% | -6.962% | 5.342% | -17.8%
| Z-Score (Absolute) |  | 0.67 | 0.69 | 2.24 | 0.84 | 0.18 | 0.76 | 0.25 | 0.94 | 0.02 | 0.87 | 0.4 | 0.68 | 0 | 0.51 | 0.08 | 0.14 | 1.21 | 0.58 | 1.18 | 0.03 | 0.54 | 0.82 | 1.73
| Analytical results for Au in OREAS 313 (Certified Value < 0.02 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.17 | 2.19 | 2.23 | 2.24 | [12 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 01g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.0027 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.0005 | < 0.001 | 0.0077 | 0.0065 | 0.015 | 0.004 | [12 empty] | 1
| 1 | 2 | 0.0023 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.0005 | < 0.001 | 0.0068 | 0.0062 | 0.017 | 0.0039 | [12 empty] | 12
| 1 | 3 | 0.0021 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.0005 | < 0.001 | 0.0071 | 0.0055 | 0.018 | 0.0034 | [12 empty] | 16
| 1 | 4 | 0.0031 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.0005 | < 0.001 | 0.0077 | 0.0071 | 0.014 | 0.0036 | [12 empty] | < 0.02
| 1 | 5 | 0.0023 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.0007 | < 0.001 | 0.0071 | 0.0062 | 0.016 | 0.0041 | [12 empty] | 69
| 1 | 6 | 0.0028 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.0008 | < 0.001 | 0.0071 | 0.006 | 0.017 | 0.0038
| Mean |  | 0.00255 | [5 empty] | 0.00075 |  | 0.00725 | 0.00625 | 0.016167 | 0.0038
| Median |  | 0.0025 | [5 empty] | 0.00075 |  | 0.0071 | 0.0062 | 0.0165 | 0.00385
| Std Dev. |  | 0.00037815 | [5 empty] | 7.0711e-05 |  | 0.00036742 | 0.00053198 | 0.001472 | 0.00026077
| Rel.Std.Dev. |  | 14.83% | [5 empty] | 9.428% |  | 5.068% | 8.512% | 9.105% | 6.862%
PDM3
| Z-Score (Absolute) |  | 2.18 | 0.49 | 0.49 | 0.49 | 0.49 | 0.49 | 2.95 | 2.92 | 0.49 | 0.85 | 2.71 | 1.74
| Analytical results for B in OREAS 313 (Certified Value < 10 ppm) | [25 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.17 | 2.19 | 2.20 | 2.24 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 1
|  |  | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | [11 empty] | 1
| 1 | 1 | 125 | < 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 6 | < 10 | < 20 | < 20 | < 10 | < 10 | [11 empty] | 1
| 1 | 2 | 130 | < 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 7 | < 10 | < 20 | < 20 | < 10 | < 10 | [11 empty] | 1
| 1 | 3 | 122 | < 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 6 | < 10 | < 20 | < 20 | < 10 | < 10 | [11 empty] | 16
| 1 | 4 | 122 | < 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 7 | < 10 | < 20 | < 20 | < 10 | < 10 | [11 empty] | < 10
| 1 | 5 | 129 | < 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 7 | < 10 | < 20 | < 20 | < 10 | < 10 | [11 empty] | 70
| 1 | 6 | 136 | < 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 7 | < 10 | < 20 | < 20 | < 10 | < 10
| Mean |  | 127.3333 | [6 empty] | 6.6667
| Median |  | 127 | [6 empty] | 7
| Std Dev. |  | 5.4283 | [6 empty] | 0.5164
| Rel.Std.Dev. |  | 4.263% | [6 empty] | 7.746%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 313 (Indicative Value 44.7 ppm) | [25 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 506 | 519.87 | 501.5 | 21.23 | 160.1145 | 20 | 20 | 10 | 50 | 30 | 5.5 | 51 | 73.563 | 30 | 39.1 | 116 | 19 | 133 | 628 | 62 | 27 |  |  |  | 1
| 1 | 2 | 507 | 512 | 472.6 | 22.25 | 150.6961 | 20 | 20 | 10 | 60 | 40 | 5.9 | 51 | 71.989 | 30 | 32.6 | 133 | 20 | 134 | 644 | 66 | 27 |  |  |  | 2
| 1 | 3 | 506 | 529 | 471.3 | 21.1 | 164.0077 | 30 | 30 | 10 | 40 | 30 | 6.5 | 54 | 66.4428 | 30 | 32.3 | 98 | 19 | 132 | 582 | 67 | 25 |  |  |  | 16
| 1 | 4 | 501 | 534.75 | 509.3 | 23.82 | 159.9619 | 20 | 30 | 20 | 50 | 30 | 4.8 | 57 | 66.8127 | 30 | 33.8 | 130 | 18 | 132 | 589 | 64 | 24 |  |  |  | 44.6718
| 1 | 5 | 507 | 515 | 487 | 23.25 | 159.7614 | 20 | 20 | 20 | 40 | 30 | 4.2 | 52 | 73.4692 | 30 | 37.4 | 109 | 19 | 132 | 586 | 71 | 25 |  |  |  | 10
| 1 | 6 | 509 | 523 | 487.9 | 24.35 | 157.3362 | 20 | 20 | < 10 | 60 | 30 | 6.4 | 54 | 63.8179 | 30 | 33.9 | 109 | 18 | 132 | 601 | 66 | 24
| Mean |  | 506 | 522.27 | 488.2667 | 22.6667 | 158.6463 | 21.6667 | 23.3333 | 14 | 50 | 31.6667 | 5.55 | 53.1667 | 69.3491 | 30 | 34.85 | 115.8333 | 18.8333 | 132.5 | 605 | 66 | 25.3333
| Median |  | 506.5 | 521.435 | 487.45 | 22.75 | 159.8617 | 20 | 20 | 10 | 50 | 30 | 5.7 | 53 | 69.4008 | 30 | 33.85 | 112.5 | 19 | 132 | 595 | 66 | 25
| Std Dev. |  | 2.6833 | 8.5517 | 15.1827 | 1.3562 | 4.4451 | 4.0825 | 5.164 | 5.4772 | 8.9443 | 4.0825 | 0.9094 | 2.3166 | 4.1755 | 0 | 2.7617 | 13.4672 | 0.75277 | 0.83666 | 25.3456 | 3.0332 | 1.3663
| Rel.Std.Dev. |  | 0.5303% | 1.637% | 3.11% | 5.983% | 2.802% | 18.84% | 22.13% | 39.12% | 17.89% | 12.89% | 16.39% | 4.357% | 6.021% | 0% | 7.925% | 11.63% | 3.997% | 0.6314% | 4.189% | 4.596% | 5.393%
| PDM3 |  | 1033% | 1069% | 993% | -49.26% | 255.1% | -51.5% | -47.77% | -68.66% | 11.93% | -29.11% | -87.58% | 19.02% | 55.24% | -32.84% | -21.99% | 159.3% | -57.84% | 196.6% | 1254% | 47.74% | -43.29%
| Z-Score (Absolute) |  | 10.85 | 11.24 | 10.43 | 0.65 | 2.59 | 0.67 | 0.63 | 0.89 | 0 | 0.44 | 1.06 | 0.08 | 0.46 | 0.48 | 0.36 | 1.57 | 0.74 | 1.96 | 13.21 | 0.38 | 0.59
| Analytical results for Be in OREAS 313 (Certified Value 0.51 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 0.6 | 0.49452 | < 2 | 0.48 | 0.51031 | 0.5 | 0.5 | 0.53 | 0.47 | 0.48 | 0.6 | 0.51 | < 0.5 | 0.36 | 0.35 | 0.6 | 0.64 | 0.59 | < 1 | [5 empty] | 1
| 1 | 2 | 0.6 | 0.43644 | < 2 | 0.57 | 0.50967 | 0.51 | 0.51 | 0.52 | 0.46 | 0.49 | 0.6 | 0.53 | < 0.5 | 0.32 | 0.36 | 0.6 | 0.65 | 0.57 | < 1 | [5 empty] | 33
| 1 | 3 | 0.6 | 0.45852 | < 2 | 0.52 | 0.51234 | 0.5 | 0.5 | 0.52 | 0.47 | 0.48 | 0.6 | 0.57 | < 0.5 | 0.25 | 0.35 | 0.6 | 0.64 | 0.55 | < 1 | [5 empty] | 16
| 1 | 4 | 0.6 | 0.5016 | < 2 | 0.55 | 0.50128 | 0.51 | 0.5 | 0.55 | 0.46 | 0.49 | 0.5 | 0.58 | < 0.5 | 0.33 | 0.36 | 0.6 | 0.61 | 0.56 | < 1 | [5 empty] | 0.5098
| 1 | 5 | 0.6 | 0.43976 | < 2 | 0.52 | 0.51265 | 0.49 | 0.5 | 0.54 | 0.46 | 0.49 | 0.7 | 0.54 | < 0.5 | 0.34 | 0.34 | 0.6 | 0.6 | 0.6 | < 1 | [5 empty] | 71
| 1 | 6 | 0.6 | 0.41664 | < 2 | 0.54 | 0.52424 | 0.5 | 0.5 | 0.53 | 0.47 | 0.48 | 0.6 | 0.49 | < 0.5 | 0.29 | 0.34 | 0.6 | 0.66 | 0.59 | < 1
| Mean |  | 0.6 | 0.45791 |  | 0.53 | 0.51175 | 0.50167 | 0.50167 | 0.53167 | 0.465 | 0.485 | 0.6 | 0.53667 |  | 0.315 | 0.35 | 0.6 | 0.63333 | 0.57667
| Median |  | 0.6 | 0.44914 |  | 0.53 | 0.51133 | 0.5 | 0.5 | 0.53 | 0.465 | 0.485 | 0.6 | 0.535 |  | 0.325 | 0.35 | 0.6 | 0.64 | 0.58
| Std Dev. |  | 0 | 0.033891 |  | 0.030984 | 0.0073902 | 0.0075277 | 0.0040825 | 0.01169 | 0.0054772 | 0.0054772 | 0.063246 | 0.034448 |  | 0.03937 | 0.0089443 | 0 | 0.023381 | 0.019664
| Rel.Std.Dev. |  | 0% | 7.401% |  | 5.846% | 1.444% | 1.501% | 0.8138% | 2.199% | 1.178% | 1.129% | 10.54% | 6.419% |  | 12.5% | 2.556% | 0% | 3.692% | 3.41%
| PDM3 |  | 17.69% | -10.18% |  | 3.962% | 0.3825% | -1.595% | -1.595% | 4.289% | -8.788% | -4.865% | 17.69% | 5.27% |  | -38.21% | -31.35% | 17.69% | 24.23% | 13.12%
| Z-Score (Absolute) |  | NA | 0.82 | 9.38 | 0.53 | 0.19 | 0 | 0 | 0.56 | 0.69 | 0.31 | NA | 0.66 | 4.74 | 3.51 | 2.85 | NA | 2.48 | 1.41 | 0.03
Results from laboratories 2.01, 2.12 and 2.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 313 (Certified Value 0.32 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 0.32 | 0.2794 | 0.37 | < 5 | 0.28 | 0.30679 | 0.35 | 0.3 | 0.31 | 0.29 | 0.31 | 0.33 | 0.31 | 0.3577 | 0.41 | 0.29 | 1.98 | 0.3 | 0.54 | 0.29 | 0.41 | 0.36 | 0.34 | 0.2 | 1
| 1 | 2 | 0.32 | 0.2626 | 0.39 | < 5 | 0.28 | 0.28978 | 0.34 | 0.32 | 0.3 | 0.32 | 0.33 | 0.34 | 0.33 | 0.34042 | 0.3 | 0.31 | 0.98 | 0.3 | 0.63 | 0.27 | 0.41 | 0.35 | 0.33 | 0.2 | 34
| 1 | 3 | 0.32 | 0.2639 | 0.38 | < 5 | 0.28 | 0.32064 | 0.37 | 0.31 | 0.31 | 0.29 | 0.33 | 0.32 | 0.32 | 0.3394 | 0.35 | 0.29 | 0.98 | 0.3 | 0.61 | 0.28 | 0.41 | 0.33 | 0.32 | 0.2 | 16
| 1 | 4 | 0.32 | 0.2614 | 0.35 | < 5 | 0.29 | 0.28625 | 0.34 | 0.32 | 0.33 | 0.32 | 0.32 | 0.32 | 0.33 | 0.33919 | 0.32 | 0.29 | 0.98 | 0.3 | 0.61 | 0.27 | 0.38 | 0.33 | 0.34 | 0.2 | 0.31715
| 1 | 5 | 0.32 | 0.2785 | 0.37 | < 5 | 0.3 | 0.31473 | 0.37 | 0.3 | 0.33 | 0.3 | 0.32 | 0.34 | 0.3 | 0.3547 | 0.3 | 0.3 | 1.98 | 0.3 | 0.63 | 0.28 | 0.41 | 0.33 | 0.33 | 0.2 | 72
| 1 | 6 | 0.32 | 0.2504 | 0.36 | < 5 | 0.31 | 0.33293 | 0.34 | 0.31 | 0.31 | 0.29 | 0.31 | 0.34 | 0.29 | 0.34393 | 0.33 | 0.28 | 0.98 | 0.3 | 0.63 | 0.28 | 0.41 | 0.35 | 0.33 | 0.2
| Mean |  | 0.32 | 0.26603 | 0.37 |  | 0.29 | 0.30852 | 0.35167 | 0.31 | 0.315 | 0.30167 | 0.32 | 0.33167 | 0.31333 | 0.34589 | 0.335 | 0.29333 | 1.3133 | 0.3 | 0.60833 | 0.27833 | 0.405 | 0.34167 | 0.33167 | 0.2
| Median |  | 0.32 | 0.26325 | 0.37 |  | 0.285 | 0.31076 | 0.345 | 0.31 | 0.31 | 0.295 | 0.32 | 0.335 | 0.315 | 0.34217 | 0.325 | 0.29 | 0.98 | 0.3 | 0.62 | 0.28 | 0.41 | 0.34 | 0.33 | 0.2
| Std Dev. |  | 0 | 0.011102 | 0.014142 |  | 0.012649 | 0.018062 | 0.01472 | 0.0089443 | 0.012247 | 0.01472 | 0.0089443 | 0.0098319 | 0.01633 | 0.0082172 | 0.041352 | 0.010328 | 0.5164 | 0 | 0.034881 | 0.0075277 | 0.012247 | 0.013292 | 0.0075277 | 3.0405e-17
| Rel.Std.Dev. |  | 0% | 4.173% | 3.822% |  | 4.362% | 5.854% | 4.186% | 2.885% | 3.888% | 4.879% | 2.795% | 2.964% | 5.212% | 2.376% | 12.34% | 3.521% | 39.32% | 0% | 5.734% | 2.705% | 3.024% | 3.89% | 2.27% | 1.52e-14%
| PDM3 |  | 0.8973% | -16.12% | 16.66% |  | -8.562% | -2.722% | 10.88% | -2.256% | -0.6792% | -4.883% | 0.8973% | 4.576% | -1.205% | 9.06% | 5.627% | -7.511% | 314.1% | -5.409% | 91.81% | -12.24% | 27.7% | 7.729% | 4.576% | -36.94%
| Z-Score (Absolute) |  | 0 | 1.68 | 1.56 | NA | 0.93 | 0.36 | 0.99 | 0.31 | 0.16 | 0.57 | 0 | 0.36 | 0.21 | 0.81 | 0.47 | 0.83 | 30.91 | NA | 8.97 | 1.3 | 2.65 | 0.67 | 0.36 | NA
Results from laboratory 2.04 were removed due to their extreme bias. Results from laboratories 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 313 (Certified Value 1.34 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.29 | 1.3394 | 1.371 | 1.28 | 1.3634 | 1.36 | 1.34 | 1.4 | 1.32 | 1.31 | 1.25 | 1.308 | 1.2943 | 1.29 | 1.29 | 1.4204 | 1.37 | 1.32 | 1.36 | 1.31 | 1.37 | 1.19 |  |  | 1
| 1 | 2 | 1.32 | 1.3466 | 1.4259 | 1.25 | 1.3581 | 1.38 | 1.38 | 1.35 | 1.3 | 1.34 | 1.26 | 1.276 | 1.3105 | 1.46 | 1.31 | 1.4117 | 1.39 | 1.29 | 1.36 | 1.3 | 1.39 | 1.15 |  |  | #N/A
| 1 | 3 | 1.32 | 1.3313 | 1.3704 | 1.22 | 1.3772 | 1.37 | 1.35 | 1.38 | 1.3 | 1.33 | 1.31 | 1.274 | 1.2916 | 1.57 | 1.31 | 1.4471 | 1.2 | 1.32 | 1.36 | 1.29 | 1.32 | 1.11 |  |  | 16
| 1 | 4 | 1.29 | 1.3229 | 1.3506 | 1.4 | 1.3679 | 1.38 | 1.37 | 1.41 | 1.31 | 1.31 | 1.28 | 1.308 | 1.3062 | 1.41 | 1.32 | 1.4283 | 1.45 | 1.32 | 1.36 | 1.29 | 1.39 | 1.15 |  |  | 1.3414
| 1 | 5 | 1.31 | 1.3527 | 1.3796 | 1.29 | 1.3741 | 1.35 | 1.36 | 1.4 | 1.32 | 1.33 | 1.5 | 1.298 | 1.2899 | 1.32 | 1.29 | 1.4428 | 1.29 | 1.3 | 1.35 | 1.29 | 1.36 | 1.14 |  |  | 73
| 1 | 6 | 1.31 | 1.3325 | 1.4161 | 1.3 | 1.4046 | 1.37 | 1.37 | 1.37 | 1.33 | 1.31 | 1.44 | 1.31 | 1.3107 | 1.47 | 1.26 | 1.402 | 1.36 | 1.29 | 1.35 | 1.31 | 1.4 | 1.14
| Mean |  | 1.3067 | 1.3375 | 1.3856 | 1.29 | 1.3742 | 1.3683 | 1.3617 | 1.385 | 1.3133 | 1.3217 | 1.34 | 1.2957 | 1.3005 | 1.42 | 1.2967 | 1.4254 | 1.3433 | 1.3067 | 1.3567 | 1.2983 | 1.3717 | 1.1467
| Median |  | 1.31 | 1.3359 | 1.3753 | 1.285 | 1.371 | 1.37 | 1.365 | 1.39 | 1.315 | 1.32 | 1.295 | 1.303 | 1.3002 | 1.435 | 1.3 | 1.4244 | 1.365 | 1.31 | 1.36 | 1.295 | 1.38 | 1.145
| Std Dev. |  | 0.013663 | 0.010915 | 0.029193 | 0.061319 | 0.016424 | 0.01169 | 0.01472 | 0.022583 | 0.012111 | 0.013292 | 0.1045 | 0.016561 | 0.0096788 | 0.10354 | 0.021602 | 0.017563 | 0.087101 | 0.015055 | 0.005164 | 0.0098319 | 0.029269 | 0.02582
| Rel.Std.Dev. |  | 1.046% | 0.816% | 2.107% | 4.753% | 1.195% | 0.8544% | 1.081% | 1.631% | 0.9221% | 1.006% | 7.798% | 1.278% | 0.7442% | 7.291% | 1.666% | 1.232% | 6.484% | 1.152% | 0.3806% | 0.7573% | 2.134% | 2.252%
| PDM3 |  | -2.588% | -0.2855% | 3.297% | -3.83% | 2.449% | 2.009% | 1.512% | 3.252% | -2.091% | -1.47% | -0.1028% | -3.408% | -3.045% | 5.861% | -3.333% | 6.263% | 0.1457% | -2.588% | 1.14% | -3.209% | 2.258% | -14.52%
| Z-Score (Absolute) |  | 0.63 | 0.02 | 0.92 | 0.96 | 0.7 | 0.58 | 0.45 | 0.91 | 0.5 | 0.34 | 0.02 | 0.85 | 0.75 | 1.6 | 0.83 | 1.71 | 0.09 | 0.63 | 0.35 | 0.8 | 0.65 | 3.79
| Analytical results for Cd in OREAS 313 (Certified Value 110 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 0
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  |  | 0
| 1 | 1 | 104 | 82.52 | 97.5 | 110.8 | 111.6962 | 110 | 116.5 | 114.5 | 108.5 | 109 | 107 | 117 | 103.7039 | 109 | 103.11 | 124 | 106 | 109 | 109.51 | 123.49 | 81.8 |  |  |  | 1
| 1 | 2 | 103 | 81.66 | 97.86 | 108.4 | 109.9848 | 111 | 119 | 111.5 | 108.5 | 111.5 | 107 | 122 | 104.7848 | 108 | 104.86 | 126.2 | 105 | 109 | 109.53 | 124.01 | 79.8 |  |  |  | 13
| 1 | 3 | 103 | 83.4075 | 97.96 | 105.8 | 112.9821 | 110.5 | 113.5 | 112.5 | 110.5 | 110 | 105 | 123 | 103.3561 | 101 | 104.46 | 105.3 | 105 | 109 | 109.28 | 118.38 | 79.9 |  |  |  | 16
| 1 | 4 | 103 | 81.975 | 97.7 | 116.3 | 112.2806 | 109.5 | 116.5 | 115 | 109 | 110.5 | 103 | 126 | 104.1947 | 107 | 105.53 | 133.1 | 104 | 106 | 108.42 | 124.53 | 79.6 |  |  |  | 109.5792
| 1 | 5 | 102 | 85.71 | 97.86 | 108.6 | 111.4305 | 110 | 115 | 116.5 | 107.5 | 110.5 | 107 | 120 | 104.533 | 114 | 103.8 | 123.7 | 106 | 110 | 105.64 | 122.67 | 81 |  |  |  | 74
| 1 | 6 | 105 | 78.36 | 97.86 | 111.1 | 112.9156 | 109.5 | 114 | 116.5 | 108.5 | 109.5 | 105 | 119 | 105.0479 | 108 | 99.73 | 125.1 | 105 | 107 | 109.31 | 125.74 | 80.8
| Mean |  | 103.3333 | 82.2721 | 97.79 | 110.1667 | 111.8816 | 110.0833 | 115.75 | 114.4167 | 108.75 | 110.1667 | 105.6667 | 121.1667 | 104.2701 | 107.8333 | 103.5817 | 122.9 | 105.1667 | 108.3333 | 108.615 | 123.1367 | 80.4833
| Median |  | 103 | 82.2475 | 97.86 | 109.7 | 111.9884 | 110 | 115.75 | 114.75 | 108.5 | 110.25 | 106 | 121 | 104.3639 | 108 | 104.13 | 124.55 | 105 | 109 | 109.295 | 123.75 | 80.35
| Std Dev. |  | 1.0328 | 2.4057 | 0.1648 | 3.5646 | 1.1203 | 0.58452 | 2.0187 | 2.0595 | 0.98742 | 0.8756 | 1.633 | 3.1885 | 0.6482 | 4.1673 | 2.0643 | 9.2885 | 0.75277 | 1.5055 | 1.5134 | 2.5475 | 0.85888
| Rel.Std.Dev. |  | 0.9995% | 2.924% | 0.1685% | 3.236% | 1.001% | 0.531% | 1.744% | 1.8% | 0.908% | 0.7948% | 1.545% | 2.632% | 0.6217% | 3.865% | 1.993% | 7.558% | 0.7158% | 1.39% | 1.393% | 2.069% | 1.067%
| PDM3 |  | -5.7% | -24.92% | -10.76% | 0.5361% | 2.101% | 0.4601% | 5.631% | 4.415% | -0.7567% | 0.5361% | -3.57% | 10.57% | -4.845% | -1.593% | -5.473% | 12.16% | -4.027% | -1.137% | -0.8799% | 12.37% | -26.55%
| Z-Score (Absolute) |  | 0.82 | 4.09 | 1.68 | 0.24 | 0.51 | 0.23 | 1.11 | 0.9 | 0.02 | 0.24 | 0.46 | 1.95 | 0.67 | 0.12 | 0.78 | 2.22 | 0.54 | 0.04 | 0 | 2.25 | 4.37
| Analytical results for Ce in OREAS 313 (Certified Value 36.3 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 48.8 | 43.179 | 50.3 | 29.339 | 38.8246 | 33.2 | 31.6 | 34.7 | 33.4 | 33.1 | 28.3 | 34.83 | 42.9844 | 15.7 | 14.49 | 41.8 | 41.8 | 46.59 | 36.41 | 24 |  |  |  |  | 1
| 1 | 2 | 48.9 | 43.187 | 51.3 | 29.019 | 38.0566 | 33.4 | 32.2 | 34.9 | 34.3 | 35.5 | 28.6 | 35.44 | 43.8439 | 15.2 | 15.15 | 41.5 | 42.6 | 44.79 | 35.99 | 24 |  |  |  |  | 14
| 1 | 3 | 48.9 | 43.171 | 50.1 | 29.104 | 39.0442 | 35 | 31.7 | 35 | 32.2 | 33.9 | 28 | 36.5 | 43.502 | 16.1 | 15.01 | 41.2 | 41.2 | 45.56 | 35.41 | 25 |  |  |  |  | 16
| 1 | 4 | 47.9 | 43.201 | 51.8 | 28.817 | 38.8423 | 34.4 | 33.2 | 34 | 34.1 | 34.6 | 27.2 | 39.46 | 43.5926 | 16 | 15.03 | 40.3 | 41.2 | 43.44 | 35.54 | 23 |  |  |  |  | 36.3202
| 1 | 5 | 47.8 | 43.233 | 50.3 | 29.946 | 39.6405 | 33.6 | 31.2 | 34.5 | 31.7 | 34.1 | 28.2 | 34.88 | 42.3133 | 16.1 | 15 | 42.3 | 41.8 | 43.74 | 36.81 | 24 |  |  |  |  | 75
| 1 | 6 | 47.4 | 43.231 | 51 | 31.109 | 40.3113 | 33.2 | 31.9 | 33.8 | 34.4 | 34.4 | 30.3 | 32.21 | 42.0062 | 16.7 | 14.62 | 41.2 | 41.7 | 46.18 | 35.53 | 25
| Mean |  | 48.2833 | 43.2003 | 50.8 | 29.5557 | 39.1199 | 33.8 | 31.9667 | 34.4833 | 33.35 | 34.2667 | 28.4333 | 35.5533 | 43.0404 | 15.9667 | 14.8833 | 41.3833 | 41.7167 | 45.05 | 35.9483 | 24.1667
| Median |  | 48.35 | 43.194 | 50.65 | 29.2215 | 38.9432 | 33.5 | 31.8 | 34.6 | 33.75 | 34.25 | 28.25 | 35.16 | 43.2432 | 16.05 | 15.005 | 41.35 | 41.75 | 45.175 | 35.765 | 24
| Std Dev. |  | 0.66156 | 0.026463 | 0.67528 | 0.85438 | 0.77344 | 0.73756 | 0.68896 | 0.48751 | 1.1502 | 0.79666 | 1.0289 | 2.3801 | 0.74364 | 0.49666 | 0.26319 | 0.67355 | 0.51543 | 1.2872 | 0.56372 | 0.75277
| Rel.Std.Dev. |  | 1.37% | 0.06126% | 1.329% | 2.891% | 1.977% | 2.182% | 2.155% | 1.414% | 3.449% | 2.325% | 3.619% | 6.695% | 1.728% | 3.111% | 1.768% | 1.628% | 1.236% | 2.857% | 1.568% | 3.115%
| PDM3 |  | 32.94% | 18.94% | 39.87% | -18.62% | 7.708% | -6.939% | -11.99% | -5.058% | -8.178% | -5.654% | -21.71% | -2.112% | 18.5% | -56.04% | -59.02% | 13.94% | 14.86% | 24.04% | -1.024% | -33.46%
| Z-Score (Absolute) |  | 1.38 | 0.85 | 1.64 | 0.57 | 0.43 | 0.13 | 0.32 | 0.06 | 0.17 | 0.08 | 0.69 | 0.06 | 0.84 | 1.98 | 2.1 | 0.66 | 0.7 | 1.04 | 0.1 | 1.13
| Analytical results for Co in OREAS 313 (Certified Value 25.8 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 22.9 | 30.5326 | 30.7 | 21.8 | 24.97 | 23.7873 | 25.3 | 26.8 | 26.3 | 26.4 | 26.3 | 25.8 | 27.2 | 25.4241 | 25.9 | 25.2 | 27.5 | 15.9 | 26.5 | 24.12 | 23.1 | 30.3 | 19.4 |  | 1
| 1 | 2 | 23 | 25.4808 | 28.8 | 21.5 | 25.81 | 22.9915 | 25.2 | 27.3 | 26.8 | 25.8 | 27.7 | 24.2 | 27.2 | 24.738 | 25.5 | 25.7 | 27.7 | 17 | 26.9 | 23.62 | 23.5 | 30.7 | 19.1 |  | 16
| 1 | 3 | 22.7 | 25.9498 | 29.6 | 21.1 | 25.13 | 23.9701 | 25.9 | 26 | 26.3 | 26.3 | 26.6 | 25.4 | 28.6 | 24.6275 | 24.1 | 25.6 | 23.3 | 16.7 | 26.5 | 23.72 | 23.2 | 29.3 | 19.1 |  | 16
| 1 | 4 | 22.7 | 27.0033 | 29.25 | 22.9 | 25.5 | 23.7649 | 24.7 | 25.2 | 27.7 | 26 | 26.9 | 24.2 | 30 | 24.3244 | 25.4 | 27 | 29.5 | 16.8 | 25.7 | 24.18 | 22.4 | 30 | 19 |  | 25.7742
| 1 | 5 | 22.9 | 26.5344 | 29.3 | 21.6 | 26.25 | 23.5999 | 26 | 25.1 | 28.2 | 25.9 | 26.8 | 25.8 | 27.1 | 24.9954 | 27.2 | 26.2 | 26.9 | 16.9 | 27.3 | 23.74 | 22.2 | 31.1 | 19.4 |  | 76
| 1 | 6 | 23 | 24.2095 | 29.2 | 21.9 | 26.62 | 24.8884 | 25.7 | 27.7 | 27.3 | 25.6 | 26.5 | 23.8 | 26.6 | 25.5444 | 25.4 | 25 | 27.4 | 16.4 | 26.5 | 23.6 | 23.5 | 32 | 19.4
| Mean |  | 22.8667 | 26.6184 | 29.475 | 21.8 | 25.7133 | 23.8337 | 25.4667 | 26.35 | 27.1 | 26 | 26.8 | 24.8667 | 27.7833 | 24.9423 | 25.5833 | 25.7833 | 27.05 | 16.6167 | 26.5667 | 23.83 | 22.9833 | 30.5667 | 19.2333
| Median |  | 22.9 | 26.2421 | 29.275 | 21.7 | 25.655 | 23.7761 | 25.5 | 26.4 | 27.05 | 25.95 | 26.7 | 24.8 | 27.2 | 24.8667 | 25.45 | 25.65 | 27.45 | 16.75 | 26.5 | 23.73 | 23.15 | 30.5 | 19.25
| Std Dev. |  | 0.13663 | 2.1457 | 0.6525 | 0.60663 | 0.64152 | 0.61676 | 0.49261 | 1.0895 | 0.77201 | 0.30332 | 0.4899 | 0.90037 | 1.275 | 0.4733 | 0.99883 | 0.72778 | 2.0413 | 0.40702 | 0.53166 | 0.25448 | 0.55648 | 0.9331 | 0.18619
| Rel.Std.Dev. |  | 0.5975% | 8.061% | 2.214% | 2.783% | 2.495% | 2.588% | 1.934% | 4.135% | 2.849% | 1.167% | 1.828% | 3.621% | 4.589% | 1.898% | 3.904% | 2.823% | 7.546% | 2.449% | 2.001% | 1.068% | 2.421% | 3.053% | 0.9681%
| PDM3 |  | -11.28% | 3.275% | 14.36% | -15.42% | -0.2362% | -7.529% | -1.193% | 2.234% | 5.144% | 0.8761% | 3.98% | -3.521% | 7.795% | -3.228% | -0.7405% | 0.03543% | 4.95% | -35.53% | 3.075% | -7.543% | -10.83% | 18.59% | -25.38%
| Z-Score (Absolute) |  | 1.44 | 0.46 | 1.9 | 1.97 | 0 | 0.95 | 0.12 | 0.32 | 0.7 | 0.14 | 0.55 | 0.43 | 1.04 | 0.39 | 0.07 | 0.04 | 0.67 | 4.59 | 0.43 | 0.95 | 1.38 | 2.45 | 3.27
| Analytical results for Cr in OREAS 313 (Certified Value 21 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 23 | 33.814 | 16.5 | 20.5 | 23.4511 | 21 | 20 | 22 | 20 | 20 | 23 | 26 | 23.4494 | 19 | 21.4 | 22 | 21 | 19 | 22.2 | 23 | 18 |  |  |  | 1
| 1 | 2 | 24 | 27.885 | 17.9 | 24 | 21.9399 | 21 | 21 | 22 | 19 | 20 | 23 | 21 | 22.5562 | 16 | 21.9 | 24 | 20 | 19 | 22.1 | 23 | 17 |  |  |  | 37
| 1 | 3 | 24 | 27.731 | 17.36 | 20.6 | 21.636 | 20 | 20 | 21 | 19 | 21 | 24 | 20 | 23.1237 | 16 | 21.6 | 20 | 20 | 19 | 21.8 | 22 | 20 |  |  |  | 16
| 1 | 4 | 24 | 28.38 | 17.5 | 20.8 | 21.4978 | 20 | 20 | 24 | 21 | 21 | 24 | 22 | 22.7589 | 18 | 21.9 | 23 | 20 | 19 | 22.4 | 22 | 18 |  |  |  | 20.9877
| 1 | 5 | 24 | 29.128 | 17.7 | 20.5 | 21.0582 | 20 | 20 | 22 | 21 | 20 | 24 | 22 | 22.0078 | 19 | 20.6 | 22 | 20 | 20 | 22.8 | 22 | 18 |  |  |  | 77
| 1 | 6 | 24 | 25.839 | 17.8 | 22.1 | 22.6066 | 20 | 20 | 21 | 19 | 20 | 24 | 24 | 21.9286 | 19 | 19.9 | 22 | 21 | 19 | 21.6 | 24 | 19
| Mean |  | 23.8333 | 28.7962 | 17.46 | 21.4167 | 22.0316 | 20.3333 | 20.1667 | 22 | 19.8333 | 20.3333 | 23.6667 | 22.5 | 22.6374 | 17.8333 | 21.2167 | 22.1667 | 20.3333 | 19.1667 | 22.15 | 22.6667 | 18.3333
| Median |  | 24 | 28.1325 | 17.6 | 20.7 | 21.788 | 20 | 20 | 22 | 19.5 | 20 | 24 | 22 | 22.6575 | 18.5 | 21.5 | 22 | 20 | 19 | 22.15 | 22.5 | 18
| Std Dev. |  | 0.40825 | 2.6897 | 0.5099 | 1.4049 | 0.8652 | 0.5164 | 0.40825 | 1.0954 | 0.98319 | 0.5164 | 0.5164 | 2.1679 | 0.60265 | 1.472 | 0.80353 | 1.3292 | 0.5164 | 0.40825 | 0.42778 | 0.8165 | 1.0328
| Rel.Std.Dev. |  | 1.713% | 9.34% | 2.92% | 6.56% | 3.927% | 2.54% | 2.024% | 4.979% | 4.957% | 2.54% | 2.182% | 9.635% | 2.662% | 8.254% | 3.787% | 5.996% | 2.54% | 2.13% | 1.931% | 3.602% | 5.633%
| PDM3 |  | 13.56% | 37.2% | -16.81% | 2.044% | 4.974% | -3.118% | -3.912% | 4.823% | -5.5% | -3.118% | 12.76% | 7.206% | 7.86% | -15.03% | 1.091% | 5.617% | -3.118% | -8.677% | 5.538% | 8% | -12.65%
| Z-Score (Absolute) |  | 1.33 | 4.08 | 2.19 | 0 | 0.34 | 0.6 | 0.69 | 0.32 | 0.87 | 0.6 | 1.24 | 0.6 | 0.67 | 1.98 | 0.11 | 0.41 | 0.6 | 1.24 | 0.41 | 0.69 | 1.7
| Analytical results for Cs in OREAS 313 (Certified Value 0.84 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | [5 empty] | 2
| 1 | 1 | 1.2 | 1.2723 | 0.82 | 1.1427 | 0.85 | 0.8 | 0.76 | 0.75 | 0.83 | 0.85 | 0.79 | 1.0114 | 0.82 | 0.52 | 0.82 | 1.11 | 1.26 | 0.98 | 0.6 | [5 empty] | 1
| 1 | 2 | 1.18 | 1.2912 | 0.85 | 1.0582 | 0.85 | 0.82 | 0.77 | 0.79 | 0.86 | 0.84 | 0.85 | 1.0152 | 0.79 | 0.56 | 0.87 | 1.12 | 1.29 | 0.95 | 0.6 | [5 empty] | 17
| 1 | 3 | 1.17 | 1.3058 | 0.8 | 1.1309 | 0.87 | 0.77 | 0.78 | 0.75 | 0.84 | 0.86 | 0.84 | 1.0282 | 0.83 | 0.55 | 0.81 | 1.08 | 1.22 | 0.92 | 0.6 | [5 empty] | 16
| 1 | 4 | 1.19 | 1.2929 | 0.74 | 1.1276 | 0.87 | 0.79 | 0.81 | 0.77 | 0.86 | 0.77 | 0.89 | 0.99987 | 0.79 | 0.55 | 0.79 | 1.1 | 1.26 | 0.93 | 0.6 | [5 empty] | 0.84491
| 1 | 5 | 1.18 | 1.3218 | 0.87 | 1.1547 | 0.86 | 0.77 | 0.82 | 0.72 | 0.85 | 0.85 | 0.79 | 1.0306 | 0.84 | 0.55 | 0.87 | 1.09 | 1.24 | 0.96 | 0.6 | [5 empty] | 78
| 1 | 6 | 1.19 | 1.2919 | 0.85 | 1.111 | 0.84 | 0.78 | 0.79 | 0.76 | 0.84 | 0.86 | 0.75 | 1.0042 | 0.77 | 0.54 | 0.86 | 1.05 | 1.25 | 1 | 0.6
| Mean |  | 1.185 | 1.296 | 0.82167 | 1.1209 | 0.85667 | 0.78833 | 0.78833 | 0.75667 | 0.84667 | 0.83833 | 0.81833 | 1.0149 | 0.80667 | 0.545 | 0.83667 | 1.0917 | 1.2533 | 0.95667 | 0.6
| Median |  | 1.185 | 1.2924 | 0.835 | 1.1292 | 0.855 | 0.785 | 0.785 | 0.755 | 0.845 | 0.85 | 0.815 | 1.0133 | 0.805 | 0.55 | 0.84 | 1.095 | 1.255 | 0.955 | 0.6
| Std Dev. |  | 0.010488 | 0.016567 | 0.047081 | 0.03403 | 0.012111 | 0.019408 | 0.023166 | 0.023381 | 0.012111 | 0.034303 | 0.050761 | 0.012454 | 0.027325 | 0.013784 | 0.034448 | 0.024833 | 0.023381 | 0.030111 | 0
| Rel.Std.Dev. |  | 0.8851% | 1.278% | 5.73% | 3.036% | 1.414% | 2.462% | 2.939% | 3.09% | 1.43% | 4.092% | 6.203% | 1.227% | 3.387% | 2.529% | 4.117% | 2.275% | 1.865% | 3.147% | 0%
| PDM3 |  | 40.25% | 53.39% | -2.751% | 32.66% | 1.391% | -6.697% | -6.697% | -10.44% | 0.2075% | -0.7788% | -3.146% | 20.12% | -4.527% | -35.5% | -0.9761% | 29.2% | 48.34% | 13.23% | -28.99%
| Z-Score (Absolute) |  | 2.86 | 3.78 | 0.14 | 2.33 | 0.15 | 0.41 | 0.41 | 0.67 | 0.07 | 0 | 0.17 | 1.46 | 0.26 | 2.42 | 0.01 | 2.09 | 3.43 | 0.98 | 1.97
| Analytical results for Cu in OREAS 313 (Certified Value 0.037 wt.%) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0328 | 0.036186 | 0.037468 | 0.03765 | 0.038373 | 0.037 | 0.0367 | 0.0386 | 0.0371 | 0.0366 | 0.0374 | 0.0363 | 0.035907 | 0.0341 | 0.038682 | 0.0392 | 0.03683 | 0.036 | 0.0361 | 0.0365 | 0.03393 | 0.03642 | 0.03476 |  | 1
| 1 | 2 | 0.0333 | 0.034563 | 0.37457 | 0.03685 | 0.037594 | 0.0375 | 0.0373 | 0.0376 | 0.0371 | 0.0368 | 0.0347 | 0.0353 | 0.036391 | 0.0356 | 0.038846 | 0.039 | 0.03684 | 0.0354 | 0.0362 | 0.0368 | 0.03338 | 0.03654 | 0.0348 |  | 39
| 1 | 3 | 0.0327 | 0.035169 | 0.03797 | 0.03516 | 0.037795 | 0.0368 | 0.0365 | 0.0382 | 0.0373 | 0.0372 | 0.0359 | 0.0352 | 0.035067 | 0.0397 | 0.039566 | 0.0398 | 0.03143 | 0.0359 | 0.0362 | 0.0356 | 0.03325 | 0.03686 | 0.03324 |  | 16
| 1 | 4 | 0.0324 | 0.035492 | 0.037456 | 0.0388 | 0.038849 | 0.0372 | 0.0372 | 0.0389 | 0.0366 | 0.0364 | 0.0352 | 0.0363 | 0.034952 | 0.0377 | 0.040164 | 0.0397 | 0.03909 | 0.0357 | 0.0363 | 0.0363 | 0.03231 | 0.0371 | 0.03378 |  | 0.036823
| 1 | 5 | 0.033 | 0.036831 | 0.03743 | 0.03843 | 0.037284 | 0.0367 | 0.0368 | 0.0382 | 0.0371 | 0.037 | 0.0365 | 0.036 | 0.035344 | 0.0369 | 0.038604 | 0.0399 | 0.03288 | 0.0356 | 0.0361 | 0.0364 | 0.0317 | 0.03646 | 0.034 |  | 79
| 1 | 6 | 0.0329 | 0.03643 | 0.03712 | 0.03786 | 0.039418 | 0.0371 | 0.0371 | 0.0382 | 0.037 | 0.0365 | 0.0355 | 0.0361 | 0.036712 | 0.038 | 0.038062 | 0.0389 | 0.03641 | 0.0351 | 0.0362 | 0.0368 | 0.03428 | 0.0368 | 0.0341
| Mean |  | 0.03285 | 0.035779 | 0.093669 | 0.037458 | 0.038219 | 0.03705 | 0.036933 | 0.038283 | 0.037033 | 0.03675 | 0.035867 | 0.035867 | 0.035729 | 0.037 | 0.038987 | 0.039417 | 0.03558 | 0.035617 | 0.036183 | 0.0364 | 0.033142 | 0.036697 | 0.034113
| Median |  | 0.03285 | 0.035839 | 0.037462 | 0.037755 | 0.038084 | 0.03705 | 0.03695 | 0.0382 | 0.0371 | 0.0367 | 0.0357 | 0.03605 | 0.035625 | 0.0373 | 0.038764 | 0.03945 | 0.03662 | 0.03565 | 0.0362 | 0.03645 | 0.033315 | 0.03667 | 0.03405
| Std Dev. |  | 0.00030166 | 0.00085182 | 0.13761 | 0.0013111 | 0.00081271 | 0.0002881 | 0.00031411 | 0.00044008 | 0.00023381 | 0.00030822 | 0.00096885 | 0.00049261 | 0.00072475 | 0.0019575 | 0.0007531 | 0.00043551 | 0.0028521 | 0.00033116 | 7.5277e-05 | 0.00044272 | 0.00097528 | 0.00026726 | 0.00059608
| Rel.Std.Dev. |  | 0.9183% | 2.381% | 146.9% | 3.5% | 2.126% | 0.7776% | 0.8505% | 1.15% | 0.6313% | 0.8387% | 2.701% | 1.373% | 2.028% | 5.291% | 1.932% | 1.105% | 8.016% | 0.9298% | 0.208% | 1.216% | 2.943% | 0.7283% | 1.747%
| PDM3 |  | -10.79% | -2.835% | 154.4% | 1.726% | 3.791% | 0.6167% | 0.2999% | 3.966% | 0.5715% | -0.198% | -2.597% | -2.597% | -2.971% | 0.4809% | 5.878% | 7.044% | -3.375% | -3.276% | -1.737% | -1.148% | -9.997% | -0.3428% | -7.358%
| Z-Score (Absolute) |  | 2.83 | 0.67 | 41.86 | 0.56 | 1.12 | 0.26 | 0.17 | 1.17 | 0.25 | 0.04 | 0.61 | 0.61 | 0.71 | 0.22 | 1.68 | 2 | 0.82 | 0.79 | 0.38 | 0.22 | 2.61 | 0 | 1.9
| Analytical results for Dy in OREAS 313 (Indicative Value 1.52 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 1.3008 | 1.694 | 1.4 | 1.79 | [20 empty] | 1
| 1 | 2 | 1.2683 | 1.6268 | 1.3 | 1.71 | [20 empty] | 5
| 1 | 3 | 1.3358 | 1.7662 | 1.3 | 1.75 | [20 empty] | 16
| 1 | 4 | 1.2789 | 1.6731 | 1.2 | 1.67 | [20 empty] | 1.5228
| 1 | 5 | 1.2927 | 1.7822 | 1.4 | 1.68 | [20 empty] | 11
| 1 | 6 | 1.2758 | 1.7534 | 1.5 | 1.8
| Mean |  | 1.292 | 1.716 | 1.35 | 1.7333
| Median |  | 1.2858 | 1.7237 | 1.35 | 1.73
| Std Dev. |  | 0.024463 | 0.060978 | 0.10488 | 0.055377
| Rel.Std.Dev. |  | 1.893% | 3.554% | 7.769% | 3.195%
| PDM3 |  | -15.15% | 12.68% | -11.35% | 13.82%
| Z-Score (Absolute) |  | 0.85 | 0.64 | 0.64 | 0.7
| Analytical results for Er in OREAS 313 (Indicative Value 0.66 ppm) | [25 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 0.619 | 0.71299 | 0.7 | 0.72 | [20 empty] | 1
| 1 | 2 | 0.595 | 0.65175 | 0.7 | 0.73 | [20 empty] | 6
| 1 | 3 | 0.5863 | 0.7193 | 0.6 | 0.75 | [20 empty] | 16
| 1 | 4 | 0.6174 | 0.64586 | 0.6 | 0.73 | [20 empty] | 0.66464
| 1 | 5 | 0.5788 | 0.70084 | 0.6 | 0.73 | [20 empty] | 12
| 1 | 6 | 0.6025 | 0.70157 | 0.6 | 0.76
| Mean |  | 0.59983 | 0.68872 | 0.63333 | 0.73667
| Median |  | 0.59875 | 0.7012 | 0.6 | 0.73
| Std Dev. |  | 0.016322 | 0.031748 | 0.05164 | 0.015055
| Rel.Std.Dev. |  | 2.721% | 4.61% | 8.154% | 2.044%
| PDM3 |  | -9.75% | 3.623% | -4.71% | 10.84%
| Z-Score (Absolute) |  | 0.93 | 0.42 | 0.42 | 1.15
| Analytical results for Eu in OREAS 313 (Indicative Value 0.62 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 0.6001 | 0.6692 | 0.5 | 0.75 | [20 empty] | 1
| 1 | 2 | 0.6167 | 0.62046 | 0.5 | 0.74 | [20 empty] | 7
| 1 | 3 | 0.6158 | 0.6379 | 0.5 | 0.74 | [20 empty] | 16
| 1 | 4 | 0.6169 | 0.63363 | 0.5 | 0.71 | [20 empty] | 0.62352
| 1 | 5 | 0.587 | 0.67797 | 0.5 | 0.69 | [20 empty] | 13
| 1 | 6 | 0.6084 | 0.69041 | 0.5 | 0.76
| Mean |  | 0.60748 | 0.65493 | 0.5 | 0.73167
| Median |  | 0.6121 | 0.65355 | 0.5 | 0.74
| Std Dev. |  | 0.011981 | 0.028019 | 0 | 0.026394
| Rel.Std.Dev. |  | 1.972% | 4.278% | 0% | 3.607%
| PDM3 |  | -2.572% | 5.037% | -19.81% | 17.34%
| Z-Score (Absolute) |  | 0.26 | 0.26 | 1.42 | 1.09
| Analytical results for Fe in OREAS 313 (Certified Value 5.91 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 5.8 | 6.0291 | 5.56 | 6.1643 | 5.92 | 5.94 | 6.24 | 5.81 | 5.84 | 5.85 | 5.91 | 6.1825 | 5.79 | 5.62 | 6.0255 | 6.14 | 5.89 | 6.05 | 5.99 | 5.89 | 5.68 |  |  |  | 1
| 1 | 2 | 5.95 | 6.0443 | 5.44 | 6.1564 | 5.98 | 6.04 | 6.09 | 5.72 | 5.88 | 5.56 | 5.77 | 6.2226 | 5.78 | 5.63 | 5.9775 | 6.14 | 5.79 | 6.06 | 5.98 | 5.96 | 5.59 |  |  |  | #N/A
| 1 | 3 | 5.93 | 6.0375 | 5.33 | 6.207 | 5.91 | 5.89 | 6.12 | 5.83 | 5.83 | 5.95 | 5.76 | 6.2868 | 5.59 | 5.65 | 6.1588 | 5.14 | 5.88 | 6.05 | 5.95 | 5.92 | 5.25 |  |  |  | 16
| 1 | 4 | 5.87 | 5.9699 | 5.96 | 6.1665 | 5.95 | 6 | 6.34 | 5.79 | 5.83 | 5.68 | 5.92 | 6.2106 | 5.75 | 5.66 | 6.1136 | 6.48 | 5.89 | 6.06 | 5.93 | 6.03 | 5.39 |  |  |  | 5.9115
| 1 | 5 | 5.87 | 6.0747 | 5.63 | 6.2445 | 5.86 | 5.91 | 6.31 | 5.78 | 5.92 | 5.89 | 5.87 | 6.3566 | 6.18 | 5.54 | 6.0863 | 5.84 | 5.82 | 6.06 | 5.93 | 5.99 | 5.41 |  |  |  | 80
| 1 | 6 | 5.86 | 6.0943 | 5.69 | 6.3219 | 5.92 | 6 | 6.18 | 5.79 | 5.82 | 5.65 | 5.88 | 6.3066 | 5.88 | 5.46 | 5.9298 | 6 | 5.77 | 6.04 | 5.99 | 5.97 | 5.42
| Mean |  | 5.88 | 6.0416 | 5.6017 | 6.2101 | 5.9233 | 5.9633 | 6.2133 | 5.7867 | 5.8533 | 5.7633 | 5.8517 | 6.2609 | 5.8283 | 5.5933 | 6.0486 | 5.9567 | 5.84 | 6.0533 | 5.9617 | 5.96 | 5.4567
| Median |  | 5.87 | 6.0409 | 5.595 | 6.1868 | 5.92 | 5.97 | 6.21 | 5.79 | 5.835 | 5.765 | 5.875 | 6.2547 | 5.785 | 5.625 | 6.0559 | 6.07 | 5.85 | 6.055 | 5.965 | 5.965 | 5.415
| Std Dev. |  | 0.053666 | 0.042885 | 0.21849 | 0.064138 | 0.040332 | 0.058878 | 0.10113 | 0.037238 | 0.038816 | 0.15462 | 0.06969 | 0.066413 | 0.19651 | 0.077889 | 0.086594 | 0.45245 | 0.053666 | 0.008165 | 0.028577 | 0.0498 | 0.15384
| Rel.Std.Dev. |  | 0.9127% | 0.7098% | 3.9% | 1.033% | 0.6809% | 0.9873% | 1.628% | 0.6435% | 0.6631% | 2.683% | 1.191% | 1.061% | 3.372% | 1.393% | 1.432% | 7.596% | 0.9189% | 0.1349% | 0.4794% | 0.8356% | 2.819%
| PDM3 |  | -0.5324% | 2.202% | -5.241% | 5.052% | 0.2006% | 0.8773% | 5.106% | -2.111% | -0.9835% | -2.506% | -1.012% | 5.912% | -1.406% | -5.382% | 2.319% | 0.7645% | -1.209% | 2.4% | 0.8491% | 0.8209% | -7.694%
| Z-Score (Absolute) |  | 0.25 | 0.67 | 1.83 | 1.63 | 0 | 0.23 | 1.65 | 0.78 | 0.4 | 0.91 | 0.41 | 1.92 | 0.54 | 1.88 | 0.71 | 0.19 | 0.48 | 0.74 | 0.22 | 0.21 | 2.66
| Analytical results for Ga in OREAS 313 (Certified Value 2.49 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 2.4 | 2.8806 | 3 | < 5 | 2.2 | 2.56 | 2.34 | 2.22 | 2.37 | 2.23 | 2.36 | 2.52 | 2.4 | 2.4437 | 4.29 | 2.6 | 2 | 1.94 | 2.4 | 3.18 | 2.98 | 2.93 | 3 |  | 1
| 1 | 2 | 2.5 | 2.9707 | 3 | < 5 | 2.2 | 2.5562 | 2.3 | 2.24 | 2.4 | 2.23 | 2.42 | 2.55 | 2.3 | 2.3714 | 4.17 | 2.6 | 2 | 1.91 | 2.3 | 3.16 | 2.96 | 2.87 | 3 |  | 21
| 1 | 3 | 2.4 | 2.9556 | 2.7 | < 5 | 2.2 | 2.42 | 2.35 | 2.24 | 2.37 | 2.22 | 2.39 | 2.66 | 2.5 | 2.3823 | 4.26 | 2.6 | 2 | 1.93 | 2.4 | 3.12 | 2.84 | 2.77 | 3 |  | 16
| 1 | 4 | 2.4 | 2.8418 | 3.1 | < 5 | 2.2 | 2.5088 | 2.33 | 2.12 | 2.43 | 2.22 | 2.36 | 2.5 | 2.5 | 2.485 | 4 | 2.8 | 2 | 1.94 | 2.3 | 3.01 | 2.99 | 2.81 | 3 |  | 2.4924
| 1 | 5 | 2.5 | 2.7536 | 3.1 | < 5 | 2.3 | 2.6345 | 2.39 | 2.22 | 2.45 | 2.17 | 2.38 | 2.68 | 2.3 | 2.4569 | 4.07 | 2.7 | 2 | 1.92 | 2.4 | 3.14 | 2.89 | 2.97 | 3 |  | 81
| 1 | 6 | 2.5 | 2.8039 | 3.1 | < 5 | 2.3 | 2.509 | 2.41 | 2.28 | 2.26 | 2.18 | 2.38 | 2.79 | 2.2 | 2.3693 | 4.07 | 2.4 | 2 | 1.9 | 2.4 | 3.25 | 2.97 | 3.12 | 3
| Mean |  | 2.45 | 2.8677 | 3 |  | 2.2333 | 2.5314 | 2.3533 | 2.22 | 2.38 | 2.2083 | 2.3817 | 2.6167 | 2.3667 | 2.4181 | 4.1433 | 2.6167 | 2 | 1.9233 | 2.3667 | 3.1433 | 2.9383 | 2.9117 | 3
| Median |  | 2.45 | 2.8612 | 3.05 |  | 2.2 | 2.5326 | 2.345 | 2.23 | 2.385 | 2.22 | 2.38 | 2.605 | 2.35 | 2.413 | 4.12 | 2.6 | 2 | 1.925 | 2.4 | 3.15 | 2.965 | 2.9 | 3
| Std Dev. |  | 0.054772 | 0.085179 | 0.15492 |  | 0.05164 | 0.071406 | 0.040332 | 0.053666 | 0.066933 | 0.026394 | 0.022286 | 0.11255 | 0.12111 | 0.049945 | 0.11587 | 0.13292 | 0 | 0.01633 | 0.05164 | 0.079162 | 0.059805 | 0.12592 | 0
| Rel.Std.Dev. |  | 2.236% | 2.97% | 5.164% |  | 2.312% | 2.821% | 1.714% | 2.417% | 2.812% | 1.195% | 0.9357% | 4.301% | 5.117% | 2.065% | 2.797% | 5.08% | 0% | 0.849% | 2.182% | 2.518% | 2.035% | 4.325% | 0%
| PDM3 |  | -1.703% | 15.06% | 20.36% |  | -10.4% | 1.564% | -5.581% | -10.93% | -4.511% | -11.4% | -4.444% | 4.984% | -5.046% | -2.983% | 66.24% | 4.984% | -19.76% | -22.83% | -5.046% | 26.11% | 17.89% | 16.82% | 20.36%
| Z-Score (Absolute) |  | 0 | 1.69 | 2.23 | 0.2 | 0.88 | 0.33 | 0.39 | 0.93 | 0.28 | 0.98 | 0.28 | 0.67 | 0.34 | 0.13 | 6.85 | 0.67 | NA | 2.13 | 0.34 | 2.81 | 1.98 | 1.87 | NA
Results from laboratories 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 313 (Indicative Value 2.6 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 3.0312 | 2.224 | 2.2 | 2.97 | [20 empty] | 1
| 1 | 2 | 3.0225 | 2.1593 | 2.1 | 3.04 | [20 empty] | 8
| 1 | 3 | 3.0149 | 2.344 | 2.1 | 2.92 | [20 empty] | 16
| 1 | 4 | 3.0415 | 2.1817 | 2.1 | 2.88 | [20 empty] | 2.5989
| 1 | 5 | 3.0755 | 2.3491 | 2.2 | 2.79 | [20 empty] | 14
| 1 | 6 | 3.0624 | 2.2586 | 2.3 | 3.01
| Mean |  | 3.0413 | 2.2528 | 2.1667 | 2.935
| Median |  | 3.0364 | 2.2413 | 2.15 | 2.945
| Std Dev. |  | 0.023526 | 0.080334 | 0.08165 | 0.091815
| Rel.Std.Dev. |  | 0.7735% | 3.566% | 3.768% | 3.128%
| PDM3 |  | 17.02% | -13.32% | -16.63% | 12.93%
| Z-Score (Absolute) |  | 0.79 | 0.6 | 0.75 | 0.6
| Analytical results for Ge in OREAS 313 (Certified Value 0.11 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [10 empty] | 2
| 1 | 1 | < 5 | < 0.2 | 0.12 | 0.09 | 0.09 | 0.11 | 0.1 | < 0.1 | < 0.1 | 0.66 | 0.14 | < 0.1 | 0.08 | 1.1 | [10 empty] | 1
| 1 | 2 | < 5 | < 0.2 | 0.11 | 0.09 | 0.09 | 0.11 | 0.11 | < 0.1 | < 0.1 | 0.64 | 0.14 | < 0.1 | 0.06 | 1.1 | [10 empty] | 23
| 1 | 3 | < 5 | < 0.2 | 0.12 | 0.11 | 0.09 | 0.1 | 0.1 | < 0.1 | < 0.1 | 0.66 | 0.13 | < 0.1 | 0.07 | 1 | [10 empty] | 16
| 1 | 4 | < 5 | < 0.2 | 0.11 | 0.1 | 0.12 | 0.11 | 0.11 | < 0.1 | < 0.1 | 0.67 | 0.13 | < 0.1 | 0.07 | 1.1 | [10 empty] | 0.11083
| 1 | 5 | < 5 | < 0.2 | 0.11 | 0.09 | 0.12 | 0.1 | 0.12 | < 0.1 | < 0.1 | 0.72 | 0.14 | < 0.1 | 0.07 | 1.1 | [10 empty] | 82
| 1 | 6 | < 5 | < 0.2 | 0.12 | 0.1 | 0.11 | 0.11 | 0.11 | < 0.1 | < 0.1 | 0.8 | 0.13 | < 0.1 | 0.07 | 1.1
| Mean |  |  |  | 0.115 | 0.096667 | 0.10333 | 0.10667 | 0.10833 |  |  | 0.69167 | 0.135 |  | 0.07 | 1.0833
| Median |  |  |  | 0.115 | 0.095 | 0.1 | 0.11 | 0.11 |  |  | 0.665 | 0.135 |  | 0.07 | 1.1
| Std Dev. |  |  |  | 0.0054772 | 0.008165 | 0.015055 | 0.005164 | 0.0075277 |  |  | 0.05947 | 0.0054772 |  | 0.0063246 | 0.040825
| Rel.Std.Dev. |  |  |  | 4.763% | 8.447% | 14.57% | 4.841% | 6.949% |  |  | 8.598% | 4.057% |  | 9.035% | 3.768%
| PDM3 |  |  |  | 3.759% | -12.78% | -6.767% | -3.759% | -2.256% |  |  | 524.1% | 21.8% |  | -36.84% | 877.4%
| Z-Score (Absolute) |  | 51.03 | 0.07 | 0.25 | 0.14 | 0 | 0.07 | 0.11 | 1.14 | 1.14 | 12.53 | 0.67 | 1.14 | 0.71 | NA
Results from laboratory 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 313 (Certified Value 0.35 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 0.39 | 0.43 | 0.34609 | 0.32 | 0.31 | 0.29 | 0.33 | 0.36 | 0.4 | 0.35 | 0.40407 | 0.49 | 0.31 | 0.49 | 0.34 | 0.4 | 0.29 | 0.3 | [6 empty] | 1
| 1 | 2 | 0.39 | 0.43 | 0.34121 | 0.33 | 0.33 | 0.31 | 0.34 | 0.36 | 0.4 | 0.35 | 0.40091 | 0.53 | 0.28 | 0.46 | 0.34 | 0.4 | 0.32 | 0.3 | [6 empty] | 30
| 1 | 3 | 0.39 | 0.4 | 0.38757 | 0.33 | 0.31 | 0.33 | 0.33 | 0.35 | 0.4 | 0.36 | 0.39507 | 0.5 | 0.3 | 0.48 | 0.31 | 0.4 | 0.3 | 0.3 | [6 empty] | 16
| 1 | 4 | 0.38 | 0.4 | 0.34491 | 0.32 | 0.34 | 0.33 | 0.34 | 0.36 | 0.4 | 0.38 | 0.41034 | 0.5 | 0.32 | 0.46 | 0.33 | 0.4 | 0.32 | 0.3 | [6 empty] | 0.35315
| 1 | 5 | 0.39 | 0.42 | 0.36855 | 0.34 | 0.34 | 0.35 | 0.33 | 0.37 | 0.4 | 0.37 | 0.39379 | 0.57 | 0.3 | 0.5 | 0.35 | 0.4 | 0.32 | 0.3 | [6 empty] | 83
| 1 | 6 | 0.38 | 0.44 | 0.34601 | 0.33 | 0.31 | 0.36 | 0.34 | 0.37 | 0.3 | 0.33 | 0.40405 | 0.55 | 0.28 | 0.46 | 0.35 | 0.42 | 0.34 | 0.3
| Mean |  | 0.38667 | 0.42 | 0.35572 | 0.32833 | 0.32333 | 0.32833 | 0.335 | 0.36167 | 0.38333 | 0.35667 | 0.40137 | 0.52333 | 0.29833 | 0.475 | 0.33667 | 0.40333 | 0.315 | 0.3
| Median |  | 0.39 | 0.425 | 0.34605 | 0.33 | 0.32 | 0.33 | 0.335 | 0.36 | 0.4 | 0.355 | 0.40248 | 0.515 | 0.3 | 0.47 | 0.34 | 0.4 | 0.32 | 0.3
| Std Dev. |  | 0.005164 | 0.016733 | 0.018404 | 0.0075277 | 0.015055 | 0.025626 | 0.0054772 | 0.0075277 | 0.040825 | 0.017512 | 0.0062007 | 0.032042 | 0.016021 | 0.017607 | 0.015055 | 0.008165 | 0.017607 | 0
| Rel.Std.Dev. |  | 1.336% | 3.984% | 5.174% | 2.293% | 4.656% | 7.805% | 1.635% | 2.081% | 10.65% | 4.91% | 1.545% | 6.123% | 5.37% | 3.707% | 4.472% | 2.024% | 5.589% | 0%
| PDM3 |  | 9.492% | 18.93% | 0.729% | -7.026% | -8.442% | -7.026% | -5.139% | 2.413% | 8.548% | 0.9967% | 13.66% | 48.19% | -15.52% | 34.5% | -4.667% | 14.21% | -10.8% | -15.05%
| Z-Score (Absolute) |  | 0.65 | 1.36 | 0.01 | 0.59 | 0.7 | 0.59 | 0.45 | 0.12 | NA | 0.01 | 0.96 | 3.56 | 1.23 | 2.53 | 0.42 | 1 | 0.88 | NA
Results from laboratories 2.12 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 313 (Certified Value 0.78 ppm) | [25 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 2.645 | 0.83125 | < 5 | 0.76 | < 1 | 0.72 | 0.82 | 0.79 | 0.76 | 0.75 | 0.78 | 0.72 | 0.73326 | 0.72 | 0.77 | 0.63 | 0.82 | 0.42 | 1.1 | 0.85 | 0.832 | 0.756 |  |  | 1
| 1 | 2 | 2.5929 | 0.853 | < 5 | 0.77 | < 1 | 0.76 | 0.69 | 0.77 | 0.75 | 0.76 | 0.76 | 0.79 | 0.74201 | 0.73 | 0.843 | 0.62 | 0.83 | 0.38 | 0.9 | 0.87 | 0.785 | 0.562 |  |  | 24
| 1 | 3 | 2.59 | 0.85951 | < 5 | 0.75 | < 1 | 0.77 | 0.82 | 0.78 | 0.72 | 0.75 | 0.77 | 0.78 | 0.72006 | 0.67 | 0.857 | 0.65 | 0.69 | 0.46 | 1 | 0.85 | 0.794 | 0.584 |  |  | 16
| 1 | 4 | 2.6063 | 0.85 | < 5 | 0.77 | < 1 | 0.78 | 0.68 | 0.83 | 0.73 | 0.73 | 0.71 | 0.82 | 0.68088 | 0.75 | 0.848 | 0.63 | 0.85 | 0.44 | 0.9 | 0.86 | 0.844 | 0.529 |  |  | 0.77914
| 1 | 5 | 2.6824 | 0.86925 | < 5 | 0.78 | < 1 | 0.76 | 0.75 | 0.86 | 0.74 | 0.76 | 0.77 | 0.74 | 0.676 | 0.79 | 0.729 | 0.66 | 0.79 | 0.42 | 1 | 0.77 | 0.78 | 0.539 |  |  | 84
| 1 | 6 | 2.5437 | 0.8816 | < 5 | 0.8 | < 1 | 0.78 | 0.74 | 0.8 | 0.74 | 0.77 | 0.76 | 0.78 | 0.71886 | 0.76 | 0.767 | 0.63 | 0.79 | 0.42 | 1 | 0.91 | 0.83 | 0.637
| Mean |  | 2.61 | 0.85743 |  | 0.77167 |  | 0.76167 | 0.75 | 0.805 | 0.74 | 0.75333 | 0.75833 | 0.77167 | 0.71184 | 0.73667 | 0.80233 | 0.63667 | 0.795 | 0.42333 | 0.98333 | 0.85167 | 0.81083 | 0.60117
| Median |  | 2.5996 | 0.85625 |  | 0.77 |  | 0.765 | 0.745 | 0.795 | 0.74 | 0.755 | 0.765 | 0.78 | 0.71946 | 0.74 | 0.8065 | 0.63 | 0.805 | 0.42 | 1 | 0.855 | 0.812 | 0.573
| Std Dev. |  | 0.048101 | 0.017254 |  | 0.017224 |  | 0.022286 | 0.060663 | 0.033912 | 0.014142 | 0.013663 | 0.024833 | 0.036009 | 0.027308 | 0.040825 | 0.053664 | 0.015055 | 0.05648 | 0.026583 | 0.075277 | 0.045789 | 0.027629 | 0.08504
| Rel.Std.Dev. |  | 1.843% | 2.012% |  | 2.232% |  | 2.926% | 8.088% | 4.213% | 1.911% | 1.814% | 3.275% | 4.666% | 3.836% | 5.542% | 6.689% | 2.365% | 7.104% | 6.279% | 7.655% | 5.376% | 3.407% | 14.15%
| PDM3 |  | 235% | 10.05% |  | -0.9593% |  | -2.243% | -3.74% | 3.319% | -5.024% | -3.312% | -2.671% | -0.9593% | -8.637% | -5.451% | 2.977% | -18.29% | 2.035% | -45.67% | 26.21% | 9.308% | 4.068% | -22.84%
| Z-Score (Absolute) |  | 28.76 | 1.49 | 27.05 | 0.16 | 4.07 | 0 | 0.18 | 0.67 | 0.34 | 0.13 | 0.05 | 0.16 | 0.78 | 0.39 | 0.63 | 1.95 | 0.52 | 5.26 | NA | 1.4 | 0.77 | 2.5
Results from laboratory 2.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ho in OREAS 313 (Indicative Value 0.27 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 0.28236 | 0.26985 | 0.2 | 0.28 | [20 empty] | 1
| 1 | 2 | 0.26767 | 0.24792 | 0.2 | 0.28 | [20 empty] | 4
| 1 | 3 | 0.27326 | 0.26441 | 0.2 | 0.28 | [20 empty] | 16
| 1 | 4 | 0.27911 | 0.2633 | 0.2 | 0.27 | [20 empty] | 0.27288
| 1 | 5 | 0.25727 | 0.28537 | 0.2 | 0.28 | [20 empty] | 10
| 1 | 6 | 0.25779 | 0.28345 | 0.2 | 0.29
| Mean |  | 0.26958 | 0.26905 | 0.2 | 0.28
| Median |  | 0.27047 | 0.26713 | 0.2 | 0.28
| Std Dev. |  | 0.010601 | 0.013968 | 3.0405e-17 | 0.0063246
| Rel.Std.Dev. |  | 3.932% | 5.191% | 1.52e-14% | 2.259%
| PDM3 |  | -1.209% | -1.402% | -26.71% | 2.611%
| Z-Score (Absolute) |  | 0.03 | 0.03 | 8.54 | 1.32
| Analytical results for In in OREAS 313 (Certified Value 0.25 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.24 | 0.25284 | < 5 | 0.26 | 0.234 | 0.263 | 0.271 | 0.262 | 0.252 | 0.253 | 0.26 | 0.272 | 0.25191 | 0.256 | 0.117 | 0.26 | 0.22 | 0.25 | 0.279 | 0.2 |  |  |  |  | 1
| 1 | 2 | 0.23 | 0.24948 | < 5 | 0.28 | 0.24443 | 0.267 | 0.285 | 0.263 | 0.256 | 0.257 | 0.24 | 0.282 | 0.24449 | 0.239 | 0.119 | 0.27 | 0.21 | 0.25 | 0.28 | 0.2 |  |  |  |  | 11
| 1 | 3 | 0.23 | 0.25354 | < 5 | 0.27 | 0.23974 | 0.281 | 0.266 | 0.265 | 0.26 | 0.26 | 0.25 | 0.276 | 0.2364 | 0.216 | 0.122 | 0.26 | 0.22 | 0.25 | 0.272 | 0.2 |  |  |  |  | 16
| 1 | 4 | 0.23 | 0.23394 | < 5 | 0.27 | 0.24027 | 0.276 | 0.267 | 0.276 | 0.265 | 0.254 | 0.23 | 0.287 | 0.22738 | 0.221 | 0.129 | 0.25 | 0.23 | 0.24 | 0.276 | 0.2 |  |  |  |  | 0.25419
| 1 | 5 | 0.24 | 0.22078 | < 5 | 0.28 | 0.26099 | 0.272 | 0.259 | 0.292 | 0.254 | 0.258 | 0.25 | 0.276 | 0.24289 | 0.234 | 0.121 | 0.27 | 0.28 | 0.25 | 0.279 | 0.2 |  |  |  |  | 85
| 1 | 6 | 0.23 | 0.23674 | < 5 | 0.27 | 0.24687 | 0.277 | 0.263 | 0.272 | 0.259 | 0.256 | 0.25 | 0.269 | 0.2312 | 0.235 | 0.111 | 0.25 | 0.24 | 0.25 | 0.285 | 0.2
| Mean |  | 0.23333 | 0.24122 |  | 0.27167 | 0.24438 | 0.27267 | 0.2685 | 0.27167 | 0.25767 | 0.25633 | 0.24667 | 0.277 | 0.23904 | 0.2335 | 0.11983 | 0.26 | 0.23333 | 0.24833 | 0.2785 | 0.2
| Median |  | 0.23 | 0.24311 |  | 0.27 | 0.24235 | 0.274 | 0.2665 | 0.2685 | 0.2575 | 0.2565 | 0.25 | 0.276 | 0.23964 | 0.2345 | 0.12 | 0.26 | 0.225 | 0.25 | 0.279 | 0.2
| Std Dev. |  | 0.005164 | 0.013007 |  | 0.0075277 | 0.0092532 | 0.0067132 | 0.0090277 | 0.011361 | 0.0046762 | 0.002582 | 0.010328 | 0.0065727 | 0.0091054 | 0.014153 | 0.005947 | 0.0089443 | 0.025033 | 0.0040825 | 0.0043243 | 3.0405e-17
| Rel.Std.Dev. |  | 2.213% | 5.392% |  | 2.771% | 3.786% | 2.462% | 3.362% | 4.182% | 1.815% | 1.007% | 4.187% | 2.373% | 3.809% | 6.061% | 4.963% | 3.44% | 10.73% | 1.644% | 1.553% | 1.52e-14%
| PDM3 |  | -8.204% | -5.101% |  | 6.877% | -3.857% | 7.27% | 5.631% | 6.877% | 1.369% | 0.8444% | -2.959% | 8.975% | -5.957% | -8.139% | -52.86% | 2.287% | -8.204% | -2.303% | 9.565% | -21.32%
| Z-Score (Absolute) |  | 1.01 | 0.66 | 98.67 | 0.67 | 0.53 | 0.72 | 0.54 | 0.67 | 0.06 | 0 | 0.43 | 0.91 | 0.76 | 1 | 6 | 0.16 | 1.01 | 0.35 | 0.97 | NA
Results from laboratory 2.25 were removed due to their 0.01 ppm reading resolution.
| Analytical results for K in OREAS 313 (Certified Value 0.245 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.346 | 0.23854 | 0.23 | 0.206 | 0.28982 | 0.24 | 0.23 | 0.26 | 0.23 | 0.23 | 0.25 | 0.244 | 0.27399 | 0.24 | 0.26 | 0.26 | 0.25 | 0.23 | 0.27 | 0.25 | 0.23 |  |  |  | 1
| 1 | 2 | 0.353 | 0.2371 | 0.22 | 0.214 | 0.28964 | 0.24 | 0.23 | 0.26 | 0.22 | 0.24 | 0.28 | 0.232 | 0.26982 | 0.25 | 0.26 | 0.25 | 0.25 | 0.23 | 0.27 | 0.25 | 0.22 |  |  |  | #N/A
| 1 | 3 | 0.346 | 0.2385 | 0.21 | 0.207 | 0.29071 | 0.24 | 0.23 | 0.26 | 0.22 | 0.23 | 0.26 | 0.233 | 0.27696 | 0.26 | 0.25 | 0.22 | 0.25 | 0.23 | 0.27 | 0.24 | 0.22 |  |  |  | 16
| 1 | 4 | 0.347 | 0.2393 | 0.24 | 0.215 | 0.28151 | 0.24 | 0.23 | 0.27 | 0.23 | 0.24 | 0.26 | 0.241 | 0.28863 | 0.25 | 0.26 | 0.26 | 0.25 | 0.23 | 0.27 | 0.25 | 0.22 |  |  |  | 0.2455
| 1 | 5 | 0.353 | 0.2291 | 0.23 | 0.215 | 0.29117 | 0.24 | 0.23 | 0.27 | 0.22 | 0.24 | 0.28 | 0.24 | 0.28477 | 0.25 | 0.24 | 0.24 | 0.24 | 0.23 | 0.27 | 0.25 | 0.22 |  |  |  | 86
| 1 | 6 | 0.338 | 0.2368 | 0.22 | 0.224 | 0.29713 | 0.24 | 0.23 | 0.26 | 0.23 | 0.23 | 0.28 | 0.228 | 0.27827 | 0.26 | 0.22 | 0.25 | 0.24 | 0.23 | 0.27 | 0.26 | 0.23
| Mean |  | 0.34717 | 0.23656 | 0.225 | 0.2135 | 0.29 | 0.24 | 0.23 | 0.26333 | 0.225 | 0.235 | 0.26833 | 0.23633 | 0.27874 | 0.25167 | 0.24833 | 0.24667 | 0.24667 | 0.23 | 0.27 | 0.25 | 0.22333
| Median |  | 0.3465 | 0.2378 | 0.225 | 0.2145 | 0.29027 | 0.24 | 0.23 | 0.26 | 0.225 | 0.235 | 0.27 | 0.2365 | 0.27762 | 0.25 | 0.255 | 0.25 | 0.25 | 0.23 | 0.27 | 0.25 | 0.22
| Std Dev. |  | 0.0055648 | 0.0037746 | 0.010488 | 0.0065345 | 0.0050007 | 0 | 0 | 0.005164 | 0.0054772 | 0.0054772 | 0.013292 | 0.0062183 | 0.0069226 | 0.0075277 | 0.016021 | 0.015055 | 0.005164 | 0 | 0 | 0.0063246 | 0.005164
| Rel.Std.Dev. |  | 1.603% | 1.596% | 4.661% | 3.061% | 1.724% | 0% | 0% | 1.961% | 2.434% | 2.331% | 4.953% | 2.631% | 2.484% | 2.991% | 6.451% | 6.104% | 2.094% | 0% | 0% | 2.53% | 2.312%
| PDM3 |  | 41.41% | -3.641% | -8.35% | -13.03% | 18.13% | -2.239% | -6.313% | 7.265% | -8.35% | -4.276% | 9.302% | -3.733% | 13.54% | 2.513% | 1.155% | 0.4761% | 0.4761% | -6.313% | 9.981% | 1.834% | -9.028%
| Z-Score (Absolute) |  | 4.07 | 0.41 | 0.88 | 1.34 | 1.75 | 0.27 | 0.67 | 0.67 | 0.88 | 0.47 | 0.88 | 0.42 | 1.3 | 0.2 | 0.07 | 0 | 0 | 0.67 | 0.94 | 0.13 | 0.94
| Analytical results for La in OREAS 313 (Certified Value 16.9 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 1
| 1 | 1 | 22.9 | 20.4788 | 17.1 | 21.8 | 12.01 | 18.2875 | 14.7 | 14.3 | 16.6 | 15 | 15.2 | 13.4 | 16.4 | 18.7867 | 5.7 | 14.4 | 14 | 12 | 16.1 | 20.9 | 22.02 | 16.1 | 10 |  | 1
| 1 | 2 | 23.8 | 20.4256 | 16.15 | 20.6 | 11.947 | 17.9999 | 14.7 | 14.8 | 15.8 | 15.8 | 16.5 | 13.5 | 16.5 | 18.8192 | 5.4 | 14.1 | 15 | 11 | 15.7 | 21 | 21.88 | 16 | 9 |  | 12
| 1 | 3 | 23.7 | 20.4408 | 15.2 | 20.5 | 12.155 | 17.9698 | 15.4 | 14.7 | 16.9 | 14.4 | 15.6 | 13.4 | 16.8 | 19.9267 | 5.8 | 13.8 | 12 | 11 | 15.6 | 20.6 | 21.97 | 15.7 | 10 |  | 16
| 1 | 4 | 23.8 | 20.5568 | 17.1 | 21.1 | 12.281 | 18.2699 | 14.9 | 14.9 | 15.4 | 15.6 | 16.1 | 12.9 | 17.7 | 19.8802 | 5.8 | 15.9 | 15 | 11 | 15.1 | 20.7 | 21.08 | 15.3 | 9 |  | 16.8537
| 1 | 5 | 23.8 | 20.5096 | 15.2 | 21.4 | 12.399 | 19.1044 | 14.4 | 13.9 | 15.8 | 14.4 | 16 | 13.6 | 15.7 | 19.7697 | 5.5 | 13.5 | 14 | 11 | 15.9 | 21 | 21.47 | 16.2 | 9 |  | 87
| 1 | 6 | 23.6 | 20.5368 | 16.15 | 21.3 | 13.068 | 19.2842 | 14.6 | 14.3 | 15.9 | 16 | 16 | 15 | 14.2 | 19.63 | 6 | 12.4 | 14 | 11 | 15.7 | 20.6 | 21.96 | 17 | 10
| Mean |  | 23.6 | 20.4914 | 16.15 | 21.1167 | 12.31 | 18.4859 | 14.7833 | 14.4833 | 16.0667 | 15.2 | 15.9 | 13.6333 | 16.2167 | 19.4687 | 5.7 | 14.0167 | 14 | 11.1667 | 15.6833 | 20.8 | 21.73 | 16.05 | 9.5
| Median |  | 23.75 | 20.4942 | 16.15 | 21.2 | 12.218 | 18.2787 | 14.7 | 14.5 | 15.85 | 15.3 | 16 | 13.45 | 16.45 | 19.6999 | 5.75 | 13.95 | 14 | 11 | 15.7 | 20.8 | 21.92 | 16.05 | 9.5
| Std Dev. |  | 0.35214 | 0.05238 | 0.84971 | 0.49565 | 0.40721 | 0.56714 | 0.34303 | 0.38166 | 0.56451 | 0.70427 | 0.44721 | 0.71181 | 1.1822 | 0.52587 | 0.21909 | 1.1514 | 1.0954 | 0.40825 | 0.33714 | 0.18974 | 0.37609 | 0.56833 | 0.54772
| Rel.Std.Dev. |  | 1.492% | 0.2556% | 5.261% | 2.347% | 3.308% | 3.068% | 2.32% | 2.635% | 3.514% | 4.633% | 2.813% | 5.221% | 7.29% | 2.701% | 3.844% | 8.214% | 7.825% | 3.656% | 2.15% | 0.9122% | 1.731% | 3.541% | 5.766%
| PDM3 |  | 40.03% | 21.58% | -4.175% | 25.29% | -26.96% | 9.685% | -12.28% | -14.06% | -4.67% | -9.812% | -5.659% | -19.11% | -3.78% | 15.52% | -66.18% | -16.83% | -16.93% | -33.74% | -6.944% | 23.41% | 28.93% | -4.769% | -43.63%
| Z-Score (Absolute) |  | 2.11 | 1.24 | 0.03 | 1.41 | 1.04 | 0.68 | 0.35 | 0.44 | 0 | 0.24 | 0.04 | 0.67 | 0.05 | 0.95 | 2.89 | 0.57 | NA | 1.36 | 0.1 | 1.33 | 1.58 | 0 | NA
Results from laboratories 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 313 (Certified Value 7 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 7.2 | 6.7146 | 7.8 | 6.32 | 7.4542 | 7.1 | 6.8 | 7.7 | 7 | 6.7 | 6.9 | 6.6 | 7.0932 | 6.3 | 4.9 | 7 | 6.89 | 7.4 | 7.1 | 5.1 |  |  |  |  | 1
| 1 | 2 | 7.2 | 6.7391 | 7.9 | 6.77 | 7.5172 | 7.2 | 6.9 | 7.4 | 6.9 | 6.9 | 6.8 | 6.3 | 7.0031 | 5.5 | 5.2 | 7 | 6.73 | 7.2 | 7.4 | 4.4 |  |  |  |  | #N/A
| 1 | 3 | 7.2 | 6.708 | 8.9 | 6.58 | 7.4957 | 7 | 6.8 | 7.4 | 7 | 6.7 | 6.9 | 6.4 | 6.9672 | 4.8 | 5.4 | 7 | 6.54 | 7.2 | 6.9 | 4.9 |  |  |  |  | 16
| 1 | 4 | 7.2 | 6.7763 | 8.4 | 7.11 | 7.4526 | 7.2 | 6.9 | 7.6 | 6.9 | 6.8 | 6.8 | 6.5 | 7.1536 | 5.9 | 5.2 | 7 | 6.86 | 7 | 6.7 | 4.1 |  |  |  |  | 7.0036
| 1 | 5 | 7.2 | 6.6955 | 7.7 | 6.77 | 7.4913 | 7 | 6.9 | 7.6 | 6.9 | 6.8 | 7 | 6.5 | 7.2323 | 6.1 | 5.2 | 7 | 7.03 | 6.8 | 7.3 | 4.5 |  |  |  |  | 88
| 1 | 6 | 7.3 | 6.6565 | 8.7 | 7.13 | 7.5941 | 7.1 | 6.9 | 7.6 | 6.9 | 6.8 | 7.2 | 6.4 | 7.1521 | 6.2 | 5 | 7 | 6.9 | 7.3 | 7.5 | 4.7
| Mean |  | 7.2167 | 6.715 | 8.2333 | 6.78 | 7.5009 | 7.1 | 6.8667 | 7.55 | 6.9333 | 6.7833 | 6.9333 | 6.45 | 7.1002 | 5.8 | 5.15 | 7 | 6.825 | 7.15 | 7.15 | 4.6167
| Median |  | 7.2 | 6.7112 | 8.15 | 6.77 | 7.4935 | 7.1 | 6.9 | 7.6 | 6.9 | 6.8 | 6.9 | 6.45 | 7.1226 | 6 | 5.2 | 7 | 6.875 | 7.2 | 7.2 | 4.6
| Std Dev. |  | 0.040825 | 0.04051 | 0.50465 | 0.31087 | 0.05212 | 0.089443 | 0.05164 | 0.12247 | 0.05164 | 0.075277 | 0.15055 | 0.10488 | 0.10015 | 0.56569 | 0.17607 | 0 | 0.16932 | 0.21679 | 0.30822 | 0.36009
| Rel.Std.Dev. |  | 0.5657% | 0.6033% | 6.129% | 4.585% | 0.6949% | 1.26% | 0.752% | 1.622% | 0.7448% | 1.11% | 2.171% | 1.626% | 1.411% | 9.753% | 3.419% | 0% | 2.481% | 3.032% | 4.311% | 7.8%
| PDM3 |  | 3.042% | -4.121% | 17.56% | -3.193% | 7.1% | 1.376% | -1.956% | 7.801% | -1.004% | -3.145% | -1.004% | -7.905% | 1.38% | -17.19% | -26.47% | -0.05181% | -2.551% | 2.09% | 2.09% | -34.08%
| Z-Score (Absolute) |  | 0.88 | 0.68 | 4.05 | 0.48 | 1.77 | 0.52 | 0.21 | 1.92 | 0 | 0.47 | 0 | 1.5 | 0.52 | 3.53 | 5.55 | NA | 0.34 | 0.67 | 0.67 | 7.21
Results from laboratory 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 313 (Indicative Value 0.092 ppm) | [25 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.04 | 2.06 | 2.12 | 2.13 | 2.21 | 2.23 | [18 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [18 empty] | 3
|  |  | 0.25g | 0.2g | 0.5g | 0.25g | 0.25g | 01g | [18 empty] | 3
| 1 | 1 | < 5 | 0.092171 | < 0.1 | 0.08 | 0.1 | 0.1 | [18 empty] | 1
| 1 | 2 | < 5 | 0.085756 | < 0.1 | 0.09 | 0.09 | 0.1 | [18 empty] | 5
| 1 | 3 | < 5 | 0.09341 | < 0.1 | 0.08 | 0.1 | 0.1 | [18 empty] | 16
| 1 | 4 | < 5 | 0.087114 | < 0.1 | 0.08 | 0.1 | 0.09 | [18 empty] | 0.091785
| 1 | 5 | < 5 | 0.093427 | < 0.1 | 0.08 | 0.1 | 0.09 | [18 empty] | 11
| 1 | 6 | < 5 | 0.090954 | < 0.1 | 0.08 | 0.1 | 0.1
| Mean |  |  | 0.090472 |  | 0.081667 | 0.098333 | 0.096667
| Median |  |  | 0.091562 |  | 0.08 | 0.1 | 0.1
| Std Dev. |  |  | 0.0032863 |  | 0.0040825 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  |  | 3.632% |  | 4.999% | 4.152% | 5.342%
| PDM3 |  |  | -1.43% |  | -11.02% | 7.135% | 5.319%
| Z-Score (Absolute) |  | 194.72 | 0.25 | 3.53 | 0.96 | 0.39 | 0.25
| Analytical results for Mg in OREAS 313 (Certified Value 0.822 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.86 | 0.83657 | 0.81 | 0.825 | 0.87983 | 0.79 | 0.79 | 0.85 | 0.77 | 0.78 | 0.81 | 0.838 | 0.82707 | 0.82 | 0.81 | 0.88 | 0.83 | 0.87 | 0.81 | 0.88 | 0.78 |  |  |  | 1
| 1 | 2 | 0.87 | 0.86827 | 0.8 | 0.843 | 0.87852 | 0.8 | 0.81 | 0.82 | 0.76 | 0.79 | 0.74 | 0.817 | 0.81545 | 0.79 | 0.83 | 0.88 | 0.83 | 0.87 | 0.8 | 0.89 | 0.78 |  |  |  | #N/A
| 1 | 3 | 0.86 | 0.84236 | 0.78 | 0.832 | 0.89174 | 0.79 | 0.79 | 0.83 | 0.77 | 0.78 | 0.78 | 0.815 | 0.80623 | 0.77 | 0.82 | 0.75 | 0.84 | 0.87 | 0.8 | 0.85 | 0.74 |  |  |  | 16
| 1 | 4 | 0.86 | 0.82363 | 0.87 | 0.809 | 0.88329 | 0.79 | 0.8 | 0.87 | 0.77 | 0.78 | 0.75 | 0.839 | 0.81024 | 0.79 | 0.84 | 0.92 | 0.83 | 0.87 | 0.8 | 0.89 | 0.76 |  |  |  | 0.8224
| 1 | 5 | 0.87 | 0.8443 | 0.82 | 0.846 | 0.89212 | 0.78 | 0.79 | 0.86 | 0.77 | 0.79 | 0.81 | 0.83 | 0.82663 | 0.85 | 0.81 | 0.83 | 0.82 | 0.88 | 0.8 | 0.87 | 0.76 |  |  |  | 89
| 1 | 6 | 0.87 | 0.86486 | 0.83 | 0.836 | 0.90321 | 0.79 | 0.8 | 0.83 | 0.77 | 0.78 | 0.78 | 0.836 | 0.8118 | 0.8 | 0.79 | 0.86 | 0.81 | 0.87 | 0.81 | 0.9 | 0.74
| Mean |  | 0.865 | 0.84667 | 0.81833 | 0.83183 | 0.88812 | 0.79 | 0.79667 | 0.84333 | 0.76833 | 0.78333 | 0.77833 | 0.82917 | 0.81624 | 0.80333 | 0.81667 | 0.85333 | 0.82667 | 0.87167 | 0.80333 | 0.88 | 0.76
| Median |  | 0.865 | 0.84333 | 0.815 | 0.834 | 0.88752 | 0.79 | 0.795 | 0.84 | 0.77 | 0.78 | 0.78 | 0.833 | 0.81362 | 0.795 | 0.815 | 0.87 | 0.83 | 0.87 | 0.8 | 0.885 | 0.76
| Std Dev. |  | 0.0054772 | 0.017055 | 0.030605 | 0.013497 | 0.0093931 | 0.0063246 | 0.008165 | 0.019664 | 0.0040825 | 0.005164 | 0.029269 | 0.010685 | 0.0087366 | 0.028048 | 0.017512 | 0.058538 | 0.010328 | 0.0040825 | 0.005164 | 0.017889 | 0.017889
| Rel.Std.Dev. |  | 0.6332% | 2.014% | 3.74% | 1.623% | 1.058% | 0.8006% | 1.025% | 2.332% | 0.5313% | 0.6592% | 3.76% | 1.289% | 1.07% | 3.491% | 2.144% | 6.86% | 1.249% | 0.4684% | 0.6428% | 2.033% | 2.354%
| PDM3 |  | 5.18% | 2.951% | -0.4942% | 1.147% | 7.992% | -3.939% | -3.129% | 2.546% | -6.574% | -4.75% | -5.358% | 0.8231% | -0.7491% | -2.318% | -0.6969% | 3.762% | 0.5191% | 5.991% | -2.318% | 7.004% | -7.587%
| Z-Score (Absolute) |  | 1.11 | 0.67 | 0 | 0.32 | 1.66 | 0.67 | 0.52 | 0.6 | 1.19 | 0.83 | 0.95 | 0.26 | 0.05 | 0.36 | 0.04 | 0.83 | 0.2 | 1.27 | 0.36 | 1.47 | 1.39
| Analytical results for Mn in OREAS 313 (Certified Value 0.185 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.168 | 0.19307 | 0.1988 | 0.1835 | 0.19449 | 0.182 | 0.183 | 0.1995 | 0.1815 | 0.1765 | 0.177 | 0.1854 | 0.18789 | 0.185 | 0.1718 | 0.1985 | 0.1481 | 0.1892 | 0.1851 | 0.18 | 0.1555 |  |  |  | 1
| 1 | 2 | 0.172 | 0.19346 | 0.2062 | 0.1809 | 0.19024 | 0.185 | 0.1875 | 0.192 | 0.1785 | 0.181 | 0.173 | 0.1808 | 0.18641 | 0.18 | 0.174 | 0.2023 | 0.1475 | 0.1887 | 0.1843 | 0.1827 | 0.1528 |  |  |  | 35
| 1 | 3 | 0.171 | 0.19342 | 0.1967 | 0.1772 | 0.19133 | 0.1845 | 0.184 | 0.1945 | 0.1805 | 0.179 | 0.187 | 0.1816 | 0.18885 | 0.175 | 0.1745 | 0.1729 | 0.1473 | 0.1878 | 0.184 | 0.1806 | 0.1529 |  |  |  | 16
| 1 | 4 | 0.164 | 0.19504 | 0.1943 | 0.1874 | 0.19316 | 0.1855 | 0.186 | 0.196 | 0.1795 | 0.178 | 0.173 | 0.1859 | 0.18426 | 0.179 | 0.1772 | 0.2143 | 0.1461 | 0.189 | 0.1831 | 0.1848 | 0.152 |  |  |  | 0.18452
| 1 | 5 | 0.169 | 0.19455 | 0.2 | 0.1866 | 0.19406 | 0.1825 | 0.184 | 0.1935 | 0.182 | 0.181 | 0.183 | 0.1838 | 0.19004 | 0.193 | 0.1728 | 0.1938 | 0.1475 | 0.1881 | 0.1802 | 0.1838 | 0.1556 |  |  |  | 90
| 1 | 6 | 0.169 | 0.19288 | 0.2058 | 0.1889 | 0.19987 | 0.184 | 0.1865 | 0.1925 | 0.183 | 0.179 | 0.182 | 0.1845 | 0.183 | 0.18 | 0.1678 | 0.198 | 0.1468 | 0.188 | 0.1849 | 0.1827 | 0.153
| Mean |  | 0.16883 | 0.19374 | 0.2003 | 0.18408 | 0.19386 | 0.18392 | 0.18517 | 0.19467 | 0.18083 | 0.17908 | 0.17917 | 0.18367 | 0.18674 | 0.182 | 0.17302 | 0.19663 | 0.14722 | 0.18847 | 0.1836 | 0.18243 | 0.15363
| Median |  | 0.169 | 0.19344 | 0.1994 | 0.18505 | 0.19361 | 0.18425 | 0.185 | 0.194 | 0.181 | 0.179 | 0.1795 | 0.18415 | 0.18715 | 0.18 | 0.1734 | 0.19825 | 0.1474 | 0.1884 | 0.18415 | 0.1827 | 0.15295
| Std Dev. |  | 0.0027869 | 0.00086151 | 0.0048241 | 0.004436 | 0.0033654 | 0.0013934 | 0.0017512 | 0.0027689 | 0.0016633 | 0.001744 | 0.0057417 | 0.0020588 | 0.0027156 | 0.006261 | 0.0031435 | 0.013571 | 0.00068823 | 0.0005785 | 0.0018111 | 0.0018381 | 0.0015267
| Rel.Std.Dev. |  | 1.651% | 0.4447% | 2.408% | 2.41% | 1.736% | 0.7576% | 0.9457% | 1.422% | 0.9198% | 0.9739% | 3.205% | 1.121% | 1.454% | 3.44% | 1.817% | 6.902% | 0.4675% | 0.307% | 0.9864% | 1.008% | 0.9937%
| PDM3 |  | -8.502% | 4.995% | 8.552% | -0.237% | 5.061% | -0.3273% | 0.3501% | 5.499% | -1.998% | -2.947% | -2.902% | -0.4628% | 1.204% | -1.366% | -6.235% | 6.564% | -20.22% | 2.139% | -0.4989% | -1.131% | -16.74%
| Z-Score (Absolute) |  | 2.18 | 1.48 | 2.45 | 0.06 | 1.5 | 0.04 | 0.22 | 1.62 | 0.42 | 0.67 | 0.66 | 0 | 0.45 | 0.25 | 1.57 | 1.91 | 5.36 | 0.71 | 0.01 | 0.18 | 4.42
| Analytical results for Mo in OREAS 313 (Certified Value 5.08 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 5 | 5.1366 | 5.12 | 6.7 | 5.28 | 5.6539 | 5.39 | 5.46 | 4.85 | 4.69 | 4.89 | 5.08 | 5.45 | 4.8779 | 4.75 | 5.15 | 5.1 | 5.1 | 3.56 | 4.77 | 5.09 | 4.8 | 5.93 | 3.7 | 1
| 1 | 2 | 5 | 5.1135 | 5.14 | 7 | 5.58 | 5.465 | 5.41 | 5.6 | 5.08 | 4.72 | 5.05 | 5.03 | 5.32 | 4.5804 | 4.65 | 5.14 | 4.9 | 5.2 | 3.61 | 4.93 | 5.28 | 5 | 5.97 | 3.6 | 14
| 1 | 3 | 5 | 5.1076 | 5.1 | 6.5 | 5.08 | 5.5026 | 5.43 | 5.45 | 4.81 | 4.71 | 5 | 4.96 | 5.4 | 4.6917 | 4.63 | 5.39 | 5.1 | 4.4 | 3.66 | 4.8 | 5.76 | 4.9 | 5.71 | 3.7 | 16
| 1 | 4 | 4.9 | 5.1743 | 5.23 | 7.1 | 5.05 | 5.7307 | 5.27 | 5.37 | 5.24 | 4.82 | 4.96 | 4.74 | 5.46 | 4.6973 | 4.56 | 5.44 | 5.2 | 5.5 | 3.64 | 4.68 | 4.78 | 4.9 | 5.91 | 3.7 | 5.0831
| 1 | 5 | 4.9 | 5.1728 | 5.12 | 6.6 | 5.32 | 5.8214 | 5.49 | 5.14 | 5.22 | 4.65 | 4.96 | 4.98 | 5.55 | 4.9004 | 4.64 | 5.32 | 5.2 | 5.1 | 3.48 | 4.91 | 4.96 | 4.9 | 5.98 | 3.7 | 91
| 1 | 6 | 4.9 | 5.1469 | 5.12 | 7.4 | 5.29 | 5.5847 | 5.49 | 5.3 | 4.73 | 4.63 | 4.97 | 4.66 | 5.33 | 5.0263 | 4.62 | 5.1 | 5.1 | 5.1 | 3.39 | 4.77 | 5.03 | 5 | 6.22 | 4
| Mean |  | 4.95 | 5.142 | 5.1383 | 6.8833 | 5.2667 | 5.6264 | 5.4133 | 5.3867 | 4.9883 | 4.7033 | 4.9717 | 4.9083 | 5.4183 | 4.7957 | 4.6417 | 5.2567 | 5.1 | 5.0667 | 3.5567 | 4.81 | 5.15 | 4.9167 | 5.9533 | 3.7333
| Median |  | 4.95 | 5.1417 | 5.12 | 6.85 | 5.285 | 5.6193 | 5.42 | 5.41 | 4.965 | 4.7 | 4.965 | 4.97 | 5.425 | 4.7876 | 4.635 | 5.235 | 5.1 | 5.1 | 3.585 | 4.785 | 5.06 | 4.9 | 5.95 | 3.7
| Std Dev. |  | 0.054772 | 0.028434 | 0.046655 | 0.34303 | 0.19159 | 0.13621 | 0.081404 | 0.15718 | 0.22049 | 0.066833 | 0.052694 | 0.16857 | 0.087044 | 0.16598 | 0.061779 | 0.14487 | 0.10954 | 0.36148 | 0.10405 | 0.094446 | 0.34059 | 0.075277 | 0.16354 | 0.13663
| Rel.Std.Dev. |  | 1.107% | 0.553% | 0.908% | 4.983% | 3.638% | 2.421% | 1.504% | 2.918% | 4.42% | 1.421% | 1.06% | 3.434% | 1.606% | 3.461% | 1.331% | 2.756% | 2.148% | 7.134% | 2.926% | 1.964% | 6.613% | 1.531% | 2.747% | 3.66%
| PDM3 |  | -2.619% | 1.158% | 1.087% | 35.42% | 3.611% | 10.69% | 6.497% | 5.972% | -1.864% | -7.471% | -2.192% | -3.438% | 6.595% | -5.654% | -8.684% | 3.415% | 0.3324% | -0.3233% | -30.03% | -5.373% | 1.316% | -3.274% | 17.12% | -26.55%
| Z-Score (Absolute) |  | 0.39 | 0.17 | 0.16 | 5.32 | 0.54 | 1.6 | 0.97 | 0.9 | 0.28 | 1.12 | 0.33 | 0.52 | 0.99 | 0.85 | 1.3 | 0.51 | 0.05 | 0.05 | 4.51 | 0.81 | 0.2 | 0.49 | 2.57 | 3.99
| Analytical results for Na in OREAS 313 (Indicative Value 0.011 wt.%) | [25 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.01 | 0.055392 | 0.01 | 0.007 | 0.01187 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.013 | 0.015 | 0.021145 | < 0.01 | 0.012 | 0.008 | 0.01 | 0.01 | 0.017 | 0.01 | 0.01 |  |  |  | 1
| 1 | 2 | 0.01 | 0.054153 | 0.01 | 0.007 | 0.011409 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.016 | 0.015 | 0.02399 | < 0.01 | 0.012 | 0.008 | 0.01 | 0.02 | 0.018 | 0.02 | 0.01 |  |  |  | 6
| 1 | 3 | 0.01 | 0.05589 | 0.01 | 0.007 | 0.011735 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.014 | 0.014 | 0.024142 | < 0.01 | 0.012 | 0.007 | 0.01 | 0.01 | 0.017 | 0.01 | 0.01 |  |  |  | 16
| 1 | 4 | 0.01 | 0.05424 | 0.01 | 0.007 | 0.011053 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.016 | 0.014 | 0.019324 | < 0.01 | 0.012 | 0.009 | 0.01 | 0.01 | 0.016 | 0.01 | < 0.01 |  |  |  | 0.011412
| 1 | 5 | 0.01 | 0.05481 | 0.01 | 0.007 | 0.011425 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.015 | 0.014 | 0.019171 | < 0.01 | 0.011 | 0.008 | 0.01 | 0.02 | 0.014 | 0.02 | 0.01 |  |  |  | 12
| 1 | 6 | 0.01 | 0.054252 | 0.01 | 0.007 | 0.011571 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.016 | 0.014 | 0.019453 | < 0.01 | 0.011 | 0.008 | 0.01 | 0.01 | 0.015 | 0.02 | 0.02
| Mean |  | 0.01 | 0.05479 | 0.01 | 0.007 | 0.01151 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.015 | 0.014333 | 0.021204 |  | 0.011667 | 0.008 | 0.01 | 0.013333 | 0.016167 | 0.015 | 0.012
| Median |  | 0.01 | 0.054531 | 0.01 | 0.007 | 0.011498 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.0155 | 0.014 | 0.020299 |  | 0.012 | 0.008 | 0.01 | 0.01 | 0.0165 | 0.015 | 0.01
| Std Dev. |  | 0 | 0.00071701 | 0 | 9.5015e-19 | 0.00028651 | 0 | 0 | 0 | 0 | 0 | 0.0012649 | 0.0005164 | 0.0023294 |  | 0.0005164 | 0.00063246 | 0 | 0.005164 | 0.001472 | 0.0054772 | 0.0044721
| Rel.Std.Dev. |  | 0% | 1.309% | 0% | 1.357e-14% | 2.489% | 0% | 0% | 0% | 0% | 0% | 8.433% | 3.603% | 10.99% |  | 4.426% | 7.906% | 0% | 38.73% | 9.105% | 36.51% | 37.27%
| PDM3 |  | -12.38% | 380.1% | -12.38% | -38.66% | 0.8594% | -12.38% | -12.38% | -12.38% | 75.25% | -12.38% | 31.44% | 25.59% | 85.8% |  | 2.228% | -29.9% | -12.38% | 16.83% | 41.66% | 31.44% | 5.149%
| Z-Score (Absolute) |  | 0.2 | 10.46 | 0.2 | 0.91 | 0.16 | 0.99 | 0.2 | 0.2 | 2.18 | 0.2 | 0.99 | 0.83 | 2.47 | 1.39 | 0.2 | 0.67 | 0.2 | 0.59 | 1.27 | 0.99 | 0
| Analytical results for Nb in OREAS 313 (Certified Value 0.057 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | < 0.2 | < 5 | 0.043074 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.2 | 0.2 | 0.46 | 0.05 | 0.07 | 0.05 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.2 | < 5 | 0.040367 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.2 | 0.09 | 0.49 | 0.06 | 0.06 | 0.06 | < 0.1 | [7 empty] | 16
| 1 | 3 | < 0.2 | < 5 | 0.042469 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.2 | 0.15 | 0.47 | 0.05 | 0.07 | 0.05 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 0.2 | < 5 | 0.04623 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.2 | 0.12 | 0.46 | 0.05 | 0.07 | < 0.05 | < 0.1 | [7 empty] | 0.056572
| 1 | 5 | < 0.2 | < 5 | 0.040898 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.06 | < 0.2 | 0.09 | 0.49 | 0.06 | 0.06 | 0.06 | < 0.1 | [7 empty] | 92
| 1 | 6 | < 0.2 | < 5 | 0.038133 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.2 | 0.14 | 0.53 | 0.05 | 0.05 | 0.06 | < 0.1
| Mean |  |  |  | 0.041862 | [6 empty] | 0.068333 |  | 0.13167 | 0.48333 | 0.053333 | 0.063333 | 0.056
| Median |  |  |  | 0.041684 | [6 empty] | 0.07 |  | 0.13 | 0.48 | 0.05 | 0.065 | 0.06
| Std Dev. |  |  |  | 0.0027563 | [6 empty] | 0.0040825 |  | 0.041673 | 0.026583 | 0.005164 | 0.008165 | 0.0054772
| Rel.Std.Dev. |  |  |  | 6.584% | [6 empty] | 5.974% |  | 31.65% | 5.5% | 9.682% | 12.89% | 9.781%
| PDM3 |  |  |  | -26% | [6 empty] | 20.79% |  | 132.7% | 754.4% | -5.725% | 11.95% | -1.012%
| Z-Score (Absolute) |  | 1.28 | 63.93 | 0.23 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.02 | 0.46 | 1.28 | 2.11 | 11.29 | 0.07 | 0.33 | 0 | 0.02
| Analytical results for Nd in OREAS 313 (Indicative Value 17.7 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 1
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 1
| 1 | 1 | 16.713 | 18.9876 | 13.3 | 22.13 | [20 empty] | 1
| 1 | 2 | 16.815 | 18.0111 | 13.4 | 21.82 | [20 empty] | 7
| 1 | 3 | 16.752 | 18.3564 | 13.3 | 21.74 | [20 empty] | 16
| 1 | 4 | 16.836 | 18.1455 | 13.1 | 21.37 | [20 empty] | 17.6571
| 1 | 5 | 16.727 | 19.5081 | 13.8 | 21.4 | [20 empty] | 13
| 1 | 6 | 16.683 | 19.4469 | 14.4 | 22.05
| Mean |  | 16.7543 | 18.7426 | 13.55 | 21.7517
| Median |  | 16.7395 | 18.672 | 13.35 | 21.78
| Std Dev. |  | 0.059832 | 0.66075 | 0.47645 | 0.31821
| Rel.Std.Dev. |  | 0.3571% | 3.525% | 3.516% | 1.463%
| PDM3 |  | -5.113% | 6.147% | -23.26% | 23.19%
| Z-Score (Absolute) |  | 0.27 | 0.27 | 1.13 | 1.08
| Analytical results for Ni in OREAS 313 (Certified Value 13.9 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 14 | 14.63 | 13.5 | 12.3 | 13.9 | 14.3556 | 14 | 13.4 | 13.9 | 12.6 | 14.3 | 13.9 | 16.2 | 14.8185 | 15 | 13.9 | 14.2 | 13 | 16.3 | 14.6 | 13.5 | 16.7 | 11.1 |  | 1
| 1 | 2 | 14 | 12.2 | 13.47 | 13.88 | 16 | 13.9202 | 13.8 | 14 | 14 | 12.6 | 14.6 | 12.9 | 15 | 13.6959 | 11.1 | 14.4 | 15.4 | 13 | 16.4 | 14.1 | 13.4 | 16.7 | 10.8 |  | 18
| 1 | 3 | 14 | 11.86 | 13.6 | 11.83 | 13.8 | 14.079 | 14.2 | 13.2 | 13.8 | 12.6 | 14.6 | 13.8 | 15.1 | 15.4094 | 11 | 14.7 | 12.8 | 13 | 16.7 | 14.3 | 13 | 16.1 | 11.7 |  | 16
| 1 | 4 | 14 | 11.55 | 13.5 | 15.77 | 14.3 | 13.7563 | 13.2 | 13.3 | 15 | 13.2 | 14.7 | 13.6 | 16.1 | 14.951 | 13.1 | 15 | 15.6 | 12 | 16 | 14.8 | 13.3 | 16.1 | 11 |  | 13.8723
| 1 | 5 | 14 | 12.51 | 13.5 | 14.13 | 13.9 | 13.6363 | 14.2 | 13.2 | 14.6 | 13.2 | 14.4 | 13.3 | 14.8 | 13.8837 | 14.3 | 15.3 | 14.4 | 13 | 16.3 | 14.8 | 12.7 | 16.9 | 11.2 |  | 93
| 1 | 6 | 14 | 10.97 | 13.2 | 12.58 | 14.6 | 14.6113 | 13.8 | 13.3 | 13.5 | 12.2 | 14.2 | 12.5 | 15.6 | 14.4486 | 12.7 | 13.7 | 14.3 | 13 | 16.2 | 14.4 | 13.3 | 17.8 | 11.5
| Mean |  | 14 | 12.2867 | 13.4617 | 13.415 | 14.4167 | 14.0598 | 13.8667 | 13.4 | 14.1333 | 12.7333 | 14.4667 | 13.3333 | 15.4667 | 14.5345 | 12.8667 | 14.5 | 14.45 | 12.8333 | 16.3167 | 14.5 | 13.2 | 16.7167 | 11.2167
| Median |  | 14 | 12.03 | 13.5 | 13.23 | 14.1 | 13.9996 | 13.9 | 13.3 | 13.95 | 12.6 | 14.5 | 13.45 | 15.35 | 14.6336 | 12.9 | 14.55 | 14.35 | 13 | 16.3 | 14.5 | 13.3 | 16.7 | 11.15
| Std Dev. |  | 0 | 1.2655 | 0.13571 | 1.4647 | 0.83287 | 0.36959 | 0.37238 | 0.30332 | 0.55737 | 0.39328 | 0.19664 | 0.5465 | 0.59217 | 0.65632 | 1.6305 | 0.6229 | 1.0035 | 0.40825 | 0.23166 | 0.28284 | 0.29665 | 0.62743 | 0.33116
| Rel.Std.Dev. |  | 0% | 10.3% | 1.008% | 10.92% | 5.777% | 2.629% | 2.685% | 2.264% | 3.944% | 3.089% | 1.359% | 4.099% | 3.829% | 4.516% | 12.67% | 4.296% | 6.945% | 3.181% | 1.42% | 1.951% | 2.247% | 3.753% | 2.952%
| PDM3 |  | 0.9203% | -11.43% | -2.96% | -3.297% | 3.924% | 1.351% | -0.04085% | -3.405% | 1.881% | -8.211% | 4.284% | -3.885% | 11.49% | 4.774% | -7.249% | 4.525% | 4.164% | -7.49% | 17.62% | 4.525% | -4.847% | 20.5% | -19.14%
| Z-Score (Absolute) |  | 0 | 1.97 | 0.62 | 0.67 | 0.48 | 0.07 | 0.15 | 0.69 | 0.15 | 1.46 | 0.54 | 0.77 | 1.69 | 0.62 | 1.31 | 0.58 | 0.52 | 1.34 | 2.67 | 0.58 | 0.92 | 3.13 | 3.21
| Analytical results for P in OREAS 313 (Certified Value 0.028 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.0272 | 0.0275 | 0.03 | 0.0276 | 0.028483 | 0.028 | 0.027 | 0.029 | 0.026 | 0.026 | 0.027 | 0.028 | 0.027551 | 0.024 | 0.027 | 0.029 | 0.0264 | 0.03 | 0.0289 | 0.0287 | 0.026 |  |  |  | 1
| 1 | 2 | 0.0279 | 0.0285 | 0.03 | 0.0288 | 0.027838 | 0.028 | 0.028 | 0.028 | 0.026 | 0.027 | 0.027 | 0.027 | 0.027633 | 0.025 | 0.027 | 0.029 | 0.0257 | 0.03 | 0.029 | 0.0291 | 0.025 |  |  |  | #N/A
| 1 | 3 | 0.0277 | 0.02739 | 0.03 | 0.0275 | 0.02834 | 0.028 | 0.027 | 0.029 | 0.026 | 0.027 | 0.027 | 0.027 | 0.027994 | 0.024 | 0.027 | 0.024 | 0.0265 | 0.03 | 0.0289 | 0.0276 | 0.025 |  |  |  | 16
| 1 | 4 | 0.0276 | 0.027 | 0.03 | 0.0289 | 0.027942 | 0.028 | 0.028 | 0.029 | 0.026 | 0.027 | 0.026 | 0.028 | 0.027901 | 0.025 | 0.028 | 0.03 | 0.0261 | 0.03 | 0.0285 | 0.0289 | 0.025 |  |  |  | 0.027695
| 1 | 5 | 0.0276 | 0.0274 | 0.03 | 0.0289 | 0.028451 | 0.027 | 0.027 | 0.029 | 0.026 | 0.027 | 0.027 | 0.028 | 0.028184 | 0.027 | 0.026 | 0.029 | 0.0267 | 0.03 | 0.0284 | 0.0289 | 0.025 |  |  |  | 94
| 1 | 6 | 0.0278 | 0.02755 | 0.03 | 0.03 | 0.029963 | 0.028 | 0.028 | 0.028 | 0.027 | 0.027 | 0.027 | 0.027 | 0.028064 | 0.026 | 0.026 | 0.029 | 0.026 | 0.03 | 0.0289 | 0.0295 | 0.026
| Mean |  | 0.027633 | 0.027557 | 0.03 | 0.028617 | 0.028503 | 0.027833 | 0.0275 | 0.028667 | 0.026167 | 0.026833 | 0.026833 | 0.0275 | 0.027888 | 0.025167 | 0.026833 | 0.028333 | 0.026233 | 0.03 | 0.028767 | 0.028783 | 0.025333
| Median |  | 0.02765 | 0.02745 | 0.03 | 0.02885 | 0.028395 | 0.028 | 0.0275 | 0.029 | 0.026 | 0.027 | 0.027 | 0.0275 | 0.027947 | 0.025 | 0.027 | 0.029 | 0.02625 | 0.03 | 0.0289 | 0.0289 | 0.025
| Std Dev. |  | 0.00024221 | 0.00050107 | 0 | 0.00093684 | 0.00076409 | 0.00040825 | 0.00054772 | 0.0005164 | 0.00040825 | 0.00040825 | 0.00040825 | 0.00054772 | 0.00024825 | 0.001169 | 0.00075277 | 0.0021602 | 0.00036697 | 0 | 0.00025033 | 0.00064005 | 0.0005164
| Rel.Std.Dev. |  | 0.8765% | 1.818% | 0% | 3.274% | 2.681% | 1.467% | 1.992% | 1.801% | 1.56% | 1.521% | 1.521% | 1.992% | 0.8902% | 4.645% | 2.805% | 7.624% | 1.399% | 0% | 0.8702% | 2.224% | 2.038%
| PDM3 |  | -0.2241% | -0.5009% | 8.321% | 3.326% | 2.915% | 0.4981% | -0.7055% | 3.507% | -5.52% | -3.113% | -3.113% | -0.7055% | 0.695% | -9.13% | -3.113% | 2.303% | -5.279% | 8.321% | 3.868% | 3.928% | -8.529%
| Z-Score (Absolute) |  | 0 | 0.06 | 1.84 | 0.76 | 0.67 | 0.16 | 0.1 | 0.8 | 1.14 | 0.62 | 0.62 | 0.1 | 0.2 | 1.91 | 0.62 | 0.54 | 1.09 | 1.84 | 0.88 | 0.89 | 1.78
| Analytical results for Pb in OREAS 313 (Certified Value 3.84 wt.%) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*OES | AR*OES | AR*OES | AR*AAS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*MS | AR*OES | AR*OES/MS | 2
|  |  | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 01g |  | 0.25g | 0.25g | 01g | 0.4g | 0.25g | 3
| 1 | 1 | > 0.5 | 3.713 | 4.0767 | 3.94 | > 0.5 | 3.9226 | 3.82 | 3.89 | 3.85 | 3.81 | 3.78 | 3.89 | 3.885 | 3.7483 | 3.37 | 3.838 | 3.6485 | 3.64 | 3.83 | > 1 | 4.05 | > 0.5 | 3.788 | > 1 | 1
| 1 | 2 | > 0.5 | 3.723 | 3.9427 | 3.95 | > 0.5 | 3.8071 | 3.83 | 3.87 | 3.85 | 3.78 | 3.78 | 3.9 | 3.889 | 3.7058 | 3.62 | 3.823 | 3.6173 | 3.62 | 3.9 | > 1 | 4 | > 0.5 | 3.73 | > 1 | 40
| 1 | 3 | > 0.5 | 3.7364 | 4.0797 | 3.92 | > 0.5 | 3.9306 | 3.84 | 3.88 | 3.81 | 3.82 | 3.78 | 3.93 | 3.805 | 3.7384 | 3.52 | 3.883 | 3.6905 | 3.74 | 3.86 | > 1 | 3.99 | > 0.5 | 3.778 | > 1 | 16
| 1 | 4 | > 0.5 | 3.714 | 4.0342 | 3.89 | > 0.5 | 3.9191 | 3.84 | 3.88 | 3.87 | 3.81 | 3.78 | 3.88 | 3.906 | 3.7319 | 3.57 | 3.84 | 3.6591 | 3.74 | 3.92 | > 1 | 3.97 | > 0.5 | 3.765 | > 1 | 3.8351
| 1 | 5 | > 0.5 | 3.7 | 3.9556 | 3.89 | > 0.5 | 3.9637 | 3.84 | 3.88 | 3.83 | 3.88 | 3.83 | 3.89 | 3.888 | 3.7174 | 3.4 | 3.749 | 3.7062 | 3.52 | 3.88 | > 1 | 4.03 | > 0.5 | 3.743 | > 1 | 95
| 1 | 6 | > 0.5 | 3.7007 | 3.9932 | 3.89 | > 0.5 | 3.9431 | 3.82 | 3.87 | 3.89 | 3.82 | 3.74 | 3.86 | 3.913 | 3.714 | 3.49 | 3.844 | 3.5871 | 3.76 | 3.85 | > 1 | 4.07 | > 0.5 | 3.709 | > 1
| Mean |  |  | 3.7145 | 4.0137 | 3.9133 |  | 3.9144 | 3.8317 | 3.8783 | 3.85 | 3.82 | 3.7817 | 3.8917 | 3.881 | 3.726 | 3.495 | 3.8295 | 3.6515 | 3.67 | 3.8733 |  | 4.0183 |  | 3.7522
| Median |  |  | 3.7135 | 4.0137 | 3.905 |  | 3.9266 | 3.835 | 3.88 | 3.85 | 3.815 | 3.78 | 3.89 | 3.8885 | 3.7247 | 3.505 | 3.839 | 3.6538 | 3.69 | 3.87 |  | 4.015 |  | 3.754
| Std Dev. |  |  | 0.013812 | 0.059306 | 0.027325 |  | 0.055004 | 0.0098319 | 0.0075277 | 0.028284 | 0.032863 | 0.028577 | 0.023166 | 0.038869 | 0.016205 | 0.096488 | 0.044221 | 0.04449 | 0.093595 | 0.033267 |  | 0.038166 |  | 0.030182
| Rel.Std.Dev. |  |  | 0.3718% | 1.478% | 0.6983% |  | 1.405% | 0.2566% | 0.1941% | 0.7347% | 0.8603% | 0.7557% | 0.5953% | 1.002% | 0.4349% | 2.761% | 1.155% | 1.218% | 2.55% | 0.8589% |  | 0.9498% |  | 0.8044%
| PDM3 |  |  | -3.145% | 4.655% | 2.039% |  | 2.066% | -0.09042% | 1.126% | 0.3876% | -0.3946% | -1.394% | 1.474% | 1.196% | -2.847% | -8.869% | -0.1469% | -4.79% | -4.306% | 0.996% |  | 4.777% |  | -2.163%
| Z-Score (Absolute) |  | NA | 0.99 | 1.54 | 0.69 | NA | 0.7 | 0 | 0.4 | 0.16 | 0.1 | 0.42 | 0.51 | 0.42 | 0.9 | 2.86 | 0.02 | 1.53 | 1.37 | 0.35 | NA | 1.58 | NA | 0.67 | NA
| Analytical results for Pd in OREAS 313 (Indicative Value 41 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [22 empty] | 1
| No | No | 2.14 | 2.23 | [22 empty] | ppb
|  |  | AR*OES/MS | AR*MS | [22 empty] | 1
|  |  | 01g | 01g | [22 empty] | 1
| 1 | 1 | < 10 | 16 | [22 empty] | 1
| 1 | 2 | < 10 | 62 | [22 empty] | 8
| 1 | 3 | < 10 | 43 | [22 empty] | 16
| 1 | 4 | < 10 | 62 | [22 empty] | 41
| 1 | 5 | < 10 | 34 | [22 empty] | 14
| 1 | 6 | < 10 | 29
| Mean |  |  | 41
| Median |  |  | 38.5
| Std Dev. |  |  | 18.4608
| Rel.Std.Dev. |  |  | 45.03%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 313 (Indicative Value 4.43 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 4.1834 | 4.783 | 3.5 | 5.54 | [20 empty] | 1
| 1 | 2 | 4.2073 | 4.5313 | 3.5 | 5.43 | [20 empty] | 19
| 1 | 3 | 4.2219 | 4.5516 | 3.5 | 5.35 | [20 empty] | 16
| 1 | 4 | 4.2117 | 4.5408 | 3.4 | 5.2 | [20 empty] | 4.4319
| 1 | 5 | 4.1928 | 4.8966 | 3.5 | 5.23 | [20 empty] | 10
| 1 | 6 | 4.2257 | 4.7596 | 3.8 | 5.43
| Mean |  | 4.2071 | 4.6772 | 3.5333 | 5.3633
| Median |  | 4.2095 | 4.6556 | 3.5 | 5.39
| Std Dev. |  | 0.016446 | 0.1561 | 0.13663 | 0.13018
| Rel.Std.Dev. |  | 0.3909% | 3.338% | 3.867% | 2.427%
| PDM3 |  | -5.072% | 5.534% | -20.28% | 21.02%
| Z-Score (Absolute) |  | 0.28 | 0.28 | 1.07 | 1.09
| Analytical results for Pt in OREAS 313 (Indicative Value < 5 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [22 empty] | 1
| No | No | 2.05 | 2.14 | [22 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [22 empty] | 2
|  |  | 0.5g | 01g | [22 empty] | 2
| 1 | 1 | 2 | < 5 | [22 empty] | 1
| 1 | 2 |  | < 5 | [22 empty] | 5
| 1 | 3 |  | < 5 | [22 empty] | 16
| 1 | 4 |  | < 5 | [22 empty] | < 5
| 1 | 5 |  | < 5 | [22 empty] | 11
| 1 | 6 | 2 | < 5
| Mean |  | 2
| Median |  | 2
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 313 (Certified Value 14.5 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | [5 empty] | 1
| 1 | 1 | 18 | 23.6 | 13.52 | 17.8447 | 13.9 | 11.8 | 15 | 13.2 | 13.8 | 14.1 | 13.97 | 15.021 | 10.9 | 7.9 | 13.6 | 14.3 | 22.88 | 16.6 | 11.9 | [5 empty] | 1
| 1 | 2 | 17.9 | 22.9 | 13.56 | 17.1136 | 13.8 | 12.1 | 14.7 | 13.4 | 14.6 | 14.9 | 13.75 | 14.6088 | 10.5 | 8.8 | 13.9 | 14.4 | 22.89 | 16.8 | 11.5 | [5 empty] | 20
| 1 | 3 | 18.2 | 22.7 | 13.78 | 17.4031 | 14.2 | 12 | 14.9 | 13 | 14.5 | 15 | 14.04 | 14.4042 | 9.7 | 8.8 | 13.7 | 14.3 | 21.83 | 16.1 | 11.7 | [5 empty] | 16
| 1 | 4 | 17.9 | 25.7 | 13.62 | 16.8537 | 14 | 11.8 | 15.4 | 13.6 | 14.4 | 13.6 | 14.55 | 14.7116 | 9.6 | 8.5 | 13.1 | 14 | 23.42 | 16.1 | 11.6 | [5 empty] | 14.5261
| 1 | 5 | 17.8 | 22.2 | 13.73 | 17.5757 | 13.9 | 11.8 | 15.3 | 13 | 14.6 | 14.9 | 13.88 | 15.0118 | 11.1 | 8.6 | 14.1 | 14.5 | 22.1 | 17.3 | 11.7 | [5 empty] | 96
| 1 | 6 | 18.1 | 23.5 | 14.39 | 18.417 | 13.6 | 11.8 | 15.7 | 13.6 | 14.5 | 15 | 13.13 | 14.509 | 10.2 | 8.4 | 13.9 | 14.3 | 23.09 | 18 | 12.1
| Mean |  | 17.9833 | 23.4333 | 13.7667 | 17.5346 | 13.9 | 11.8833 | 15.1667 | 13.3 | 14.4 | 14.5833 | 13.8867 | 14.7111 | 10.3333 | 8.5 | 13.7167 | 14.3 | 22.7017 | 16.8167 | 11.75
| Median |  | 17.95 | 23.2 | 13.675 | 17.4894 | 13.9 | 11.8 | 15.15 | 13.3 | 14.5 | 14.9 | 13.925 | 14.6602 | 10.35 | 8.55 | 13.8 | 14.3 | 22.885 | 16.7 | 11.7
| Std Dev. |  | 0.1472 | 1.2258 | 0.32098 | 0.554 | 0.2 | 0.13292 | 0.36697 | 0.27568 | 0.30332 | 0.59133 | 0.46081 | 0.25767 | 0.61536 | 0.33466 | 0.34881 | 0.16733 | 0.60924 | 0.73598 | 0.21679
| Rel.Std.Dev. |  | 0.8185% | 5.231% | 2.332% | 3.159% | 1.439% | 1.119% | 2.42% | 2.073% | 2.106% | 4.055% | 3.318% | 1.752% | 5.955% | 3.937% | 2.543% | 1.17% | 2.684% | 4.376% | 1.845%
| PDM3 |  | 23.8% | 61.32% | -5.228% | 20.71% | -4.31% | -18.19% | 4.41% | -8.44% | -0.8679% | 0.3942% | -4.402% | 1.273% | -28.86% | -41.48% | -5.572% | -1.556% | 56.28% | 15.77% | -19.11%
| Z-Score (Absolute) |  | 2.48 | 6.16 | 0.36 | 2.18 | 0.27 | 1.63 | 0.58 | 0.67 | 0.07 | 0.19 | 0.28 | 0.28 | 2.67 | 3.91 | 0.39 | 0 | 5.67 | 1.7 | 1.72
| Analytical results for Re in OREAS 313 (Indicative Value 0.003 ppm) | [25 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 3
|  |  | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | [8 empty] | 3
| 1 | 1 | < 0.05 | 0.001 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | < 0.001 | 0.001 | 0.03 |  | 0.001 | < 0.005 | [8 empty] | 1
| 1 | 2 | < 0.05 |  | < 0.01 | 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.02 | 0.001 | < 0.001 | < 0.005 | [8 empty] | 27
| 1 | 3 | < 0.05 |  | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.001 | 0.03 |  | < 0.001 | < 0.005 | [8 empty] | 16
| 1 | 4 | < 0.05 |  | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.001 | 0.02 | 0.001 | < 0.001 | < 0.005 | [8 empty] | 0.0025972
| 1 | 5 | < 0.05 |  | < 0.01 | 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.001 | 0.03 |  | < 0.001 | < 0.005 | [8 empty] | 12
| 1 | 6 | < 0.05 |  | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.001 | 0.03 |  | < 0.001 | < 0.005
| Mean |  |  | 0.001 |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.0011667 |  | 0.001 | 0.001 | 0.028 | 0.001 | 0.001
| Median |  |  | 0.001 |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 |  | 0.001 | 0.001 | 0.03 | 0.001 | 0.001
| Std Dev. |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0.00040825 |  | 0 |  | 0.0044721 | 0
| Rel.Std.Dev. |  |  |  |  | 0% | 0% | 0% | 0% | 0% | 34.99% |  | 0% |  | 15.97% | 0%
| PDM3 |  |  | -61.5% |  | -61.5% | -61.5% | -61.5% | -61.5% | -61.5% | -55.08% |  | -61.5% | -61.5% | 978.1% | -61.5% | -61.5%
| Z-Score (Absolute) |  | 64.73 | 0 | 10.79 | 0 | 0.45 | 0 | 0 | 0 | 0.45 | 64.73 | 0.9 | 1.12 | 64.73 | 0 | 1.12 | 4.05
| Analytical results for S in OREAS 313 (Certified Value 4.55 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 4.36 | 5.072 | 4.7406 | 1.26 | 4.5 | 4.6059 | 4.48 | 4.56 | 4.65 | 4.37 | 4.41 | 4.378 | 4.65 | 4.4525 | 4.24 | 4.51 | 4.5575 | 4.6 | 4.43 | 4.72 | 4.34 | 4.58 | 4.99 |  | 1
| 1 | 2 | 4.47 | 4.97 | 4.5428 | 1.59 | 4.78 | 4.5916 | 4.53 | 4.64 | 4.51 | 4.35 | 4.47 | 4.328 | 4.53 | 4.4837 | 4.44 | 4.61 | 4.5377 | 4.65 | 4.46 | 4.78 | 4.32 | 4.66 | 4.78 |  | #N/A
| 1 | 3 | 4.51 | 5.018 | 4.711 | 1.47 | 4.58 | 4.6438 | 4.5 | 4.56 | 4.57 | 4.34 | 4.44 | 4.497 | 4.5 | 4.4452 | 4.79 | 4.59 | 4.6259 | 4.01 | 4.46 | 4.71 | 4.3 | 4.61 | 4.73 |  | 16
| 1 | 4 | 4.38 | 5.0012 | 4.5205 | 1.59 | 4.52 | 4.5695 | 4.53 | 4.64 | 4.67 | 4.34 | 4.41 | 4.237 | 4.63 | 4.4638 | 4.58 | 4.69 | 4.5743 | 4.97 | 4.55 | 4.79 | 4.28 | 4.71 | 4.79 |  | 4.5493
| 1 | 5 | 4.48 | 5.0397 | 4.4736 | 1.65 | 4.64 | 4.6498 | 4.45 | 4.54 | 4.59 | 4.35 | 4.46 | 4.484 | 4.59 | 4.515 | 4.54 | 4.47 | 4.6546 | 4.45 | 4.5 | 4.77 | 4.27 | 4.66 | 4.87 |  | 97
| 1 | 6 | 4.45 | 4.9851 | 4.4883 | 1.72 | 4.8 | 4.7217 | 4.49 | 4.6 | 4.57 | 4.38 | 4.42 | 3.455 | 4.6 | 4.5184 | 4.65 | 4.4 | 4.4416 | 4.58 | 4.49 | 4.77 | 4.32 | 4.65 | 4.88
| Mean |  | 4.4417 | 5.0143 | 4.5795 | 1.5467 | 4.6367 | 4.6304 | 4.4967 | 4.59 | 4.5933 | 4.355 | 4.435 | 4.2298 | 4.5833 | 4.4798 | 4.54 | 4.545 | 4.5653 | 4.5433 | 4.4817 | 4.7567 | 4.305 | 4.645 | 4.84
| Median |  | 4.46 | 5.0096 | 4.5317 | 1.59 | 4.61 | 4.6249 | 4.495 | 4.58 | 4.58 | 4.35 | 4.43 | 4.353 | 4.595 | 4.4738 | 4.56 | 4.55 | 4.5659 | 4.59 | 4.475 | 4.77 | 4.31 | 4.655 | 4.83
| Std Dev. |  | 0.059133 | 0.037354 | 0.11628 | 0.16281 | 0.12863 | 0.054231 | 0.030768 | 0.043359 | 0.058538 | 0.016432 | 0.025884 | 0.39191 | 0.05785 | 0.031437 | 0.18772 | 0.10502 | 0.074635 | 0.31341 | 0.041673 | 0.033267 | 0.026646 | 0.045056 | 0.092952
| Rel.Std.Dev. |  | 1.331% | 0.7449% | 2.539% | 10.53% | 2.774% | 1.171% | 0.6842% | 0.9446% | 1.274% | 0.3773% | 0.5836% | 9.265% | 1.262% | 0.7018% | 4.135% | 2.311% | 1.635% | 6.898% | 0.9299% | 0.6994% | 0.619% | 0.97% | 1.92%
| PDM3 |  | -2.365% | 10.22% | 0.6638% | -66% | 1.921% | 1.783% | -1.156% | 0.8953% | 0.9685% | -4.27% | -2.512% | -7.022% | 0.7487% | -1.528% | -0.2038% | -0.0939% | 0.3516% | -0.1305% | -1.486% | 4.559% | -5.369% | 2.104% | 6.391%
| Z-Score (Absolute) |  | 0.82 | 3.71 | 0.27 | 23.68 | 0.72 | 0.67 | 0.38 | 0.36 | 0.38 | 1.5 | 0.87 | 2.49 | 0.3 | 0.52 | 0.04 | 0 | 0.16 | 0.01 | 0.5 | 1.67 | 1.9 | 0.79 | 2.33
| Analytical results for Sb in OREAS 313 (Certified Value 33.1 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 2
| 1 | 1 | 31.8 | 34.84 | 40.03 | 40.9 | 30.78 | 37.8535 | 33.4 | 34.5 | 25.6 | 31.7 | 34.7 | 35.7 | 31.87 | 29.7358 | 36 | 35.33 | 17.12 | 28.2 | 22.22 | 32.3 | 31.63 | 34.41 | 36.03 | 28.83 | 1
| 1 | 2 | 32.7 | 30.28 | 41.67 | 42.4 | 31.02 | 36.654 | 32.6 | 35.2 | 26.2 | 32.9 | 35.4 | 33.8 | 32.54 | 33.0755 | 33.8 | 35.4 | 16.77 | 31.8 | 22.29 | 33.2 | 31.57 | 34.59 | 35.56 | 28.73 | 21
| 1 | 3 | 32.7 | 31.74 | 40.31 | 40.1 | 29.45 | 37.5814 | 35 | 33.1 | 27.2 | 33 | 35.6 | 35 | 32.58 | 30.2558 | 34.3 | 36.25 | 17.96 | 24.2 | 22.96 | 32.6 | 30.83 | 33.93 | 33.84 | 28.96 | 16
| 1 | 4 | 32.5 | 32.71 | 40.53 | 40.6 | 29.68 | 38.2654 | 33.4 | 34.6 | 26.3 | 32.9 | 34.7 | 33.6 | 34.52 | 33.0619 | 32.7 | 36.48 | 16.66 | 31.4 | 24.07 | 32.5 | 30.7 | 33.62 | 34.53 | 28.6 | 33.0633
| 1 | 5 | 32.8 | 32.66 | 40.17 | 43 | 31.82 | 38.5131 | 33.5 | 31.9 | 25.7 | 31.4 | 35.1 | 34.5 | 31.74 | 33.34 | 37.2 | 37.72 | 17.69 | 28.7 | 21.97 | 33.2 | 30.69 | 33.37 | 35.42 | 28.5 | 98
| 1 | 6 | 32.7 | 32.05 | 40.5 | 42 | 30.97 | 38.4953 | 33.1 | 33.1 | 24.6 | 31.9 | 34.8 | 33.3 | 30.68 | 32.4065 | 36.3 | 36.05 | 16.87 | 28.9 | 21.14 | 32.5 | 30.85 | 33.95 | 36.36 | 28.03
| Mean |  | 32.5333 | 32.38 | 40.535 | 41.5 | 30.62 | 37.8938 | 33.5 | 33.7333 | 25.9333 | 32.3 | 35.05 | 34.3167 | 32.3217 | 31.9792 | 35.05 | 36.205 | 17.1783 | 28.8667 | 22.4417 | 32.7167 | 31.045 | 33.9783 | 35.29 | 28.6083
| Median |  | 32.7 | 32.355 | 40.405 | 41.45 | 30.875 | 38.0595 | 33.4 | 33.8 | 25.95 | 32.4 | 34.95 | 34.15 | 32.205 | 32.7342 | 35.15 | 36.15 | 16.995 | 28.8 | 22.255 | 32.55 | 30.84 | 33.94 | 35.49 | 28.665
| Std Dev. |  | 0.37238 | 1.4945 | 0.58794 | 1.1349 | 0.89449 | 0.7099 | 0.80498 | 1.2372 | 0.86641 | 0.71274 | 0.38341 | 0.9196 | 1.2794 | 1.5756 | 1.7167 | 0.87299 | 0.53041 | 2.7303 | 0.99087 | 0.38687 | 0.43524 | 0.46089 | 0.94471 | 0.32677
| Rel.Std.Dev. |  | 1.145% | 4.615% | 1.45% | 2.735% | 2.921% | 1.873% | 2.403% | 3.668% | 3.341% | 2.207% | 1.094% | 2.68% | 3.958% | 4.927% | 4.898% | 2.411% | 3.088% | 9.458% | 4.415% | 1.182% | 1.402% | 1.356% | 2.677% | 1.142%
| PDM3 |  | -1.603% | -2.067% | 22.6% | 25.52% | -7.39% | 14.61% | 1.321% | 2.027% | -21.56% | -2.309% | 6.009% | 3.791% | -2.243% | -3.279% | 6.009% | 9.502% | -48.04% | -12.69% | -32.13% | -1.048% | -6.104% | 2.768% | 6.735% | -13.47%
| Z-Score (Absolute) |  | 0.03 | 0.07 | 2.41 | 2.7 | 0.61 | 1.6 | 0.27 | 0.34 | 2.04 | 0.1 | 0.74 | 0.51 | 0.09 | 0.2 | 0.74 | 1.09 | 4.7 | 1.14 | 3.1 | 0.03 | 0.48 | 0.41 | 0.81 | 1.22
| Analytical results for Sc in OREAS 313 (Certified Value 1.93 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2 | 1.9228 | 1.9 | 1.5 | 1.9 | 1.9828 | 2.1 | 1.8 | 2 | 2 | 2 | 2 | 1.9 | 1.8835 | 1.9 | 1.9 | 1.8 | 1.3 | 2.6 | 1.82 | 2.1 | 2.1 | 1.7 |  | 1
| 1 | 2 | 2 | 1.9936 | 1.7 | 1.4 | 1.9 | 1.9121 | 2.1 | 1.8 | 2 | 1.9 | 2.1 | 1.9 | 2 | 1.8999 | 1.8 | 1.9 | 1.8 | 1.4 | 2.5 | 1.93 | 2 | 2.1 | 1.6 |  | 22
| 1 | 3 | 2 | 1.9226 | 1.7 | 1.4 | 1.87 | 1.9927 | 2.2 | 1.8 | 2 | 1.9 | 2 | 2 | 2.1 | 1.7976 | 1.9 | 2 | 1.5 | 1.3 | 2.6 | 1.95 | 2 | 2 | 1.7 |  | 16
| 1 | 4 | 2 | 1.8803 | 1.8 | 1.6 | 1.86 | 1.9331 | 2 | 1.8 | 2.1 | 2 | 2 | 1.7 | 2.1 | 1.8949 | 1.9 | 2.1 | 1.9 | 1.3 | 2.4 | 1.84 | 2.1 | 2.1 | 1.7 |  | 1.9286
| 1 | 5 | 2 | 1.8472 | 1.8 | 1.5 | 1.96 | 1.9888 | 2.1 | 1.8 | 2.1 | 1.9 | 2 | 1.9 | 1.9 | 1.8881 | 1.8 | 2 | 1.8 | 1.4 | 2.6 | 1.88 | 1.9 | 2.2 | 1.7 |  | 99
| 1 | 6 | 2 | 1.9207 | 1.7 | 1.5 | 1.93 | 2.0122 | 2.2 | 1.8 | 2.1 | 1.9 | 2 | 2.1 | 1.9 | 1.797 | 2 | 1.8 | 1.8 | 1.3 | 2.5 | 1.73 | 2.2 | 2.3 | 1.8
| Mean |  | 2 | 1.9145 | 1.7667 | 1.4833 | 1.9033 | 1.9703 | 2.1167 | 1.8 | 2.05 | 1.9333 | 2.0167 | 1.9333 | 1.9833 | 1.8602 | 1.8833 | 1.95 | 1.7667 | 1.3333 | 2.5333 | 1.8583 | 2.05 | 2.1333 | 1.7
| Median |  | 2 | 1.9216 | 1.75 | 1.5 | 1.9 | 1.9858 | 2.1 | 1.8 | 2.05 | 1.9 | 2 | 1.95 | 1.95 | 1.8858 | 1.9 | 1.95 | 1.8 | 1.3 | 2.55 | 1.86 | 2.05 | 2.1 | 1.7
| Std Dev. |  | 0 | 0.049267 | 0.08165 | 0.075277 | 0.037238 | 0.038819 | 0.075277 | 0 | 0.054772 | 0.05164 | 0.040825 | 0.13663 | 0.098319 | 0.049045 | 0.075277 | 0.10488 | 0.13663 | 0.05164 | 0.08165 | 0.080353 | 0.10488 | 0.10328 | 0.063246
| Rel.Std.Dev. |  | 0% | 2.573% | 4.622% | 5.075% | 1.956% | 1.97% | 3.556% | 0% | 2.672% | 2.671% | 2.024% | 7.067% | 4.957% | 2.637% | 3.997% | 5.379% | 7.734% | 3.873% | 3.223% | 4.324% | 5.116% | 4.841% | 3.72%
| PDM3 |  | 3.702% | -0.7302% | -8.396% | -23.09% | -1.31% | 2.163% | 9.752% | -6.668% | 6.295% | 0.2457% | 4.567% | 0.2457% | 2.838% | -3.549% | -2.347% | 1.11% | -8.396% | -30.87% | 31.36% | -3.643% | 6.295% | 10.62% | -11.85%
| Z-Score (Absolute) |  | NA | 0.06 | 0.98 | 2.74 | 0.13 | 0.29 | 1.2 | 0.77 | 0.78 | 0.06 | 0.58 | 0.06 | 0.37 | 0.4 | 0.25 | 0.16 | 0.98 | 3.68 | 3.79 | 0.41 | 0.78 | 1.3 | 1.39
Results from laboratory 2.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 313 (Certified Value 1.94 ppm) | [25 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 2 | 1.4123 | 1.7 | < 5 | < 1.5 | 3.4 | 3.2 | 2.9 | 1.7 | 2.7 | 2.7 | 2 | 1.8148 | 0.6 | 1.7 | 1.92 | 1.8 | < 5 | 5 | 2 | 1.7 | 2.4 |  |  | 1
| 1 | 2 | 2 | 1.4531 | 1.7 | < 5 | < 1.5 | 3.3 | 3.1 | 2.7 | 1.8 | 2.8 | 2.7 | 2 | 1.7604 | 0.7 | 1.3 | 1.83 | 1.8 | < 5 | 4 | 2 | 1.6 | 2.4 |  |  | 2
| 1 | 3 | 2 | 1.4743 | 1.7 | < 5 | < 1.5 | 3.5 | 2.6 | 2.6 | 1.5 | 2.9 | 2.6 | 2 | 1.815 | 0.6 | 1.3 | 1.93 | 1.7 | < 5 | 5 | 2 | 1.6 | 2.7 |  |  | 16
| 1 | 4 | 2 | 1.4531 | 1.9 | < 5 | < 1.5 | 3.5 | 3.2 | 4 | 1.7 | 2.5 | 2.6 | 2 | 1.7484 | 0.4 | 1.7 | 2.04 | 2.1 | < 5 | 4 | 2 | 1.7 | 2.3 |  |  | 1.9408
| 1 | 5 | 2 | 1.3682 | 1.5 | < 5 | < 1.5 | 3.4 | 3.6 | 4.1 | 1.7 | 3 | 2.6 | 2 | 1.7169 | 0.6 | 1.8 | 1.95 | 1.9 | < 5 | 4 | 2 | 1.7 | 2.2 |  |  | 100
| 1 | 6 | 2 | 1.4743 | 1.7 | < 5 | < 1.5 | 3.4 | 3.3 | 4.6 | 1.6 | 2.7 | 2.8 | 2 | 1.7791 | 0.7 | 1.7 | 1.87 | 1.8 | < 5 | 5 | 2 | 1.8 | 2
| Mean |  | 2 | 1.4392 | 1.7 |  |  | 3.4167 | 3.1667 | 3.4833 | 1.6667 | 2.7667 | 2.6667 | 2 | 1.7724 | 0.6 | 1.5833 | 1.9233 | 1.85 |  | 4.5 | 2 | 1.6833 | 2.3333
| Median |  | 2 | 1.4531 | 1.7 |  |  | 3.4 | 3.2 | 3.45 | 1.7 | 2.75 | 2.65 | 2 | 1.7697 | 0.6 | 1.7 | 1.925 | 1.8 |  | 4.5 | 2 | 1.7 | 2.35
| Std Dev. |  | 0 | 0.041509 | 0.12649 |  |  | 0.075277 | 0.3266 | 0.85186 | 0.10328 | 0.17512 | 0.08165 | 0 | 0.038599 | 0.10954 | 0.22286 | 0.072019 | 0.13784 |  | 0.54772 | 0 | 0.075277 | 0.23381
| Rel.Std.Dev. |  | 0% | 2.884% | 7.441% |  |  | 2.203% | 10.31% | 24.46% | 6.197% | 6.33% | 3.062% | 0% | 2.178% | 18.26% | 14.08% | 3.744% | 7.451% |  | 12.17% | 0% | 4.472% | 10.02%
| PDM3 |  | 3.048% | -25.85% | -12.41% |  |  | 76.04% | 63.16% | 79.48% | -14.13% | 42.55% | 37.4% | 3.048% | -8.676% | -69.09% | -18.42% | -0.9018% | -4.68% |  | 131.9% | 3.048% | -13.27% | 20.22%
| Z-Score (Absolute) |  | NA | 0.57 | 0.21 | 0.91 | 1.53 | 2.19 | 1.84 | 2.28 | 0.26 | 1.28 | 1.14 | NA | 0.11 | 1.74 | 0.37 | 0.1 | 0 | NA | NA | NA | 0.23 | 0.67
Results from laboratories 2.01, 2.13, 2.20, 2.21 and 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 313 (Indicative Value 3.15 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 2.02 | 2.06 | 2.12 | 2.23 | [20 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 0.25g | 0.2g | 0.5g | 01g | [20 empty] | 2
| 1 | 1 | 2.9594 | 3.0234 | 2.4 | 4.15 | [20 empty] | 1
| 1 | 2 | 2.9825 | 2.8686 | 2.5 | 4.24 | [20 empty] | 24
| 1 | 3 | 2.9645 | 3.1407 | 2.8 | 4.03 | [20 empty] | 16
| 1 | 4 | 2.9882 | 2.8524 | 2.3 | 3.96 | [20 empty] | 3.1512
| 1 | 5 | 2.9829 | 2.9062 | 2.7 | 3.99 | [20 empty] | 13
| 1 | 6 | 2.9576 | 3.0631 | 2.7 | 4.17
| Mean |  | 2.9725 | 2.9757 | 2.5667 | 4.09
| Median |  | 2.9735 | 2.9648 | 2.6 | 4.09
| Std Dev. |  | 0.013508 | 0.11718 | 0.19664 | 0.11225
| Rel.Std.Dev. |  | 0.4544% | 3.938% | 7.661% | 2.744%
| PDM3 |  | -5.671% | -5.57% | -18.55% | 29.79%
| Z-Score (Absolute) |  | 0.01 | 0.01 | 1.34 | 3.68
| Analytical results for Sn in OREAS 313 (Certified Value 0.52 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | [5 empty] | 2
| 1 | 1 | 0.5 | < 5 | 0.47 | 0.42923 | 0.5 | 0.5 | 0.5 | 0.4 | 0.4 | 0.56 | 0.5 | 0.53525 | 3.4 | 2.2 | 0.4 | 0.61 | 0.63 | 0.6 | 0.2 | [5 empty] | 1
| 1 | 2 | 0.6 | < 5 | 0.49 | 0.44133 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.6 | 0.5 | 0.50597 | 0.4 | 2.2 | 0.5 | 0.61 | 0.58 | 0.6 | 0.2 | [5 empty] | 31
| 1 | 3 | 0.6 | < 5 | 0.47 | 0.40527 | 0.5 | 0.5 | 0.5 | 0.4 | 0.5 | 0.59 | 0.5 | 0.53047 | 0.4 | 2.1 | 0.5 | 0.59 | 0.59 | 0.5 | 0.2 | [5 empty] | 16
| 1 | 4 | 0.6 | < 5 | 0.46 | 0.45912 | 0.5 | 0.5 | 0.7 | 0.5 | 0.5 | 0.59 | 0.5 | < 0.5 | 0.5 | 2.1 | 0.4 | 0.48 | 0.63 | 0.6 | 0.2 | [5 empty] | 0.51544
| 1 | 5 | 0.5 | < 5 | 0.48 | 0.46852 | 0.5 | 0.4 | 0.7 | 0.4 | 0.5 | 0.57 | 0.5 | < 0.5 | 0.5 | 2.2 | 0.5 | 0.51 | 0.6 | 0.6 | 0.2 | [5 empty] | 101
| 1 | 6 | 0.6 | < 5 | 0.46 | 0.4553 | 0.5 | 0.5 | 0.5 | 0.4 | 0.5 | 0.61 | 0.5 | < 0.5 | 0.4 | 2.4 | 0.5 | 0.64 | 0.62 | 0.7 | 0.2
| Mean |  | 0.56667 |  | 0.47167 | 0.44313 | 0.5 | 0.48333 | 0.56667 | 0.43333 | 0.48333 | 0.58667 | 0.5 | 0.5239 | 0.93333 | 2.2 | 0.46667 | 0.57333 | 0.60833 | 0.6 | 0.2
| Median |  | 0.6 |  | 0.47 | 0.44831 | 0.5 | 0.5 | 0.5 | 0.4 | 0.5 | 0.59 | 0.5 | 0.53047 | 0.45 | 2.2 | 0.5 | 0.6 | 0.61 | 0.6 | 0.2
| Std Dev. |  | 0.05164 |  | 0.01169 | 0.023145 | 0 | 0.040825 | 0.10328 | 0.05164 | 0.040825 | 0.018619 | 0 | 0.015707 | 1.2094 | 0.10954 | 0.05164 | 0.063456 | 0.02137 | 0.063246 | 3.0405e-17
| Rel.Std.Dev. |  | 9.113% |  | 2.479% | 5.223% | 0% | 8.447% | 18.23% | 11.92% | 8.447% | 3.174% | 0% | 2.998% | 129.6% | 4.979% | 11.07% | 11.07% | 3.513% | 10.54% | 1.52e-14%
| PDM3 |  | 9.939% |  | -8.492% | -14.03% | -2.995% | -6.229% | 9.939% | -15.93% | -6.229% | 13.82% | -2.995% | 1.641% | 81.08% | 326.8% | -9.462% | 11.23% | 18.02% | 16.41% | -61.2%
| Z-Score (Absolute) |  | 0.67 | 20.23 | 0.29 | 0.58 | 0 | 0.17 | 0.67 | 0.67 | 0.17 | 0.88 | 0 | 1.14 | 4.38 | 17.19 | 0.34 | 0.74 | 1.1 | 1.01 | 3.03
| Analytical results for Sr in OREAS 313 (Certified Value 18.6 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 1
| 1 | 1 | 19.5 | 25.01 | 16.2 | 24.2 | 15.72 | 23.6952 | 17.6 | 15.6 | 18.4 | 16.9 | 18.2 | 15.1 | 17.6 | 19.0877 | 9.6 | 17.3 | 17 | 10 | 15.2 | 18.6 | 21.97 | 20 | 17 |  | 1
| 1 | 2 | 19.3 | 21.2 | 16.8 | 22.9 | 16.86 | 23.531 | 17.6 | 15.2 | 18.3 | 17.7 | 18.8 | 14.7 | 16.8 | 17.3263 | 9.5 | 17.3 | 18 | 10.2 | 15.2 | 18 | 21.91 | 20 | 16 |  | 32
| 1 | 3 | 19.6 | 21.48 | 16.5 | 22.6 | 16.07 | 23.0442 | 18.2 | 15.2 | 18.4 | 17.4 | 18.5 | 15.5 | 16.9 | 19.0565 | 9.7 | 16.9 | 15 | 10.2 | 15.2 | 18.5 | 21.24 | 19.3 | 15 |  | 16
| 1 | 4 | 19.4 | 22.22 | 16.4 | 24.8 | 16.34 | 22.8616 | 17.5 | 15.4 | 18.8 | 17.4 | 18.6 | 13.8 | 17.6 | 17.8641 | 9.2 | 18.4 | 19 | 10.4 | 14.9 | 18 | 21.74 | 19 | 15 |  | 18.6217
| 1 | 5 | 19.5 | 22.41 | 16.3 | 24 | 16.74 | 24.2976 | 17.8 | 15.2 | 19 | 17.2 | 18.4 | 14.8 | 17.4 | 18.5508 | 9.3 | 16.5 | 18 | 10.2 | 15.5 | 18 | 21.22 | 20.9 | 16 |  | 102
| 1 | 6 | 19.2 | 20.35 | 17.3 | 24.4 | 16.86 | 23.6342 | 17.9 | 15.2 | 18 | 17.7 | 18.1 | 15.8 | 16.7 | 17.9499 | 10.1 | 15.6 | 18 | 10.2 | 15.1 | 18 | 22.01 | 21.7 | 16
| Mean |  | 19.4167 | 22.1117 | 16.5833 | 23.8167 | 16.4317 | 23.5106 | 17.7667 | 15.3 | 18.4833 | 17.3833 | 18.4333 | 14.95 | 17.1667 | 18.3059 | 9.5667 | 17 | 17.5 | 10.2 | 15.1833 | 18.1833 | 21.6817 | 20.15 | 15.8333
| Median |  | 19.45 | 21.85 | 16.45 | 24.1 | 16.54 | 23.5826 | 17.7 | 15.2 | 18.4 | 17.4 | 18.45 | 14.95 | 17.15 | 18.2503 | 9.55 | 17.1 | 18 | 10.2 | 15.2 | 18 | 21.825 | 20 | 16
| Std Dev. |  | 0.1472 | 1.6022 | 0.40702 | 0.87274 | 0.47068 | 0.51146 | 0.2582 | 0.16733 | 0.36009 | 0.30605 | 0.2582 | 0.70071 | 0.41312 | 0.70938 | 0.32042 | 0.93381 | 1.3784 | 0.12649 | 0.19408 | 0.28577 | 0.36185 | 1.0055 | 0.75277
| Rel.Std.Dev. |  | 0.7581% | 7.246% | 2.454% | 3.664% | 2.864% | 2.175% | 1.453% | 1.094% | 1.948% | 1.761% | 1.401% | 4.687% | 2.407% | 3.875% | 3.349% | 5.493% | 7.877% | 1.24% | 1.278% | 1.572% | 1.669% | 4.99% | 4.754%
| PDM3 |  | 4.269% | 18.74% | -10.95% | 27.9% | -11.76% | 26.25% | -4.592% | -17.84% | -0.7433% | -6.65% | -1.012% | -19.72% | -7.814% | -1.696% | -48.63% | -8.709% | -6.024% | -45.23% | -18.46% | -2.354% | 16.43% | 8.207% | -14.97%
| Z-Score (Absolute) |  | 0.52 | 1.5 | NA | 2.12 | 0.56 | 2.01 | 0.08 | 0.97 | 0.18 | 0.21 | 0.17 | 1.1 | 0.29 | 0.12 | 3.05 | 0.35 | NA | 2.82 | 1.01 | 0.08 | 1.34 | 0.79 | NA
| Analytical results for Ta in OREAS 313 (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 3
|  |  | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.04 | 0.02 | < 0.05 | < 0.3 |  | < 0.01 | < 0.1 | [6 empty] | 1
| 1 | 2 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.01 | < 0.05 | 0.03 | < 0.01 | < 0.05 | < 0.3 |  | < 0.01 | < 0.1 | [6 empty] | 11
| 1 | 3 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.01 | < 0.05 | < 0.3 |  | < 0.01 | < 0.1 | [6 empty] | 16
| 1 | 4 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.01 | < 0.05 | < 0.3 |  | < 0.01 | < 0.1 | [6 empty] | < 0.01
| 1 | 5 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.02 | < 0.01 | < 0.05 | < 0.3 |  | < 0.01 | < 0.1 | [6 empty] | 103
| 1 | 6 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | < 0.05 | 0.03 | < 0.01 | < 0.05 | < 0.3 | 0.01 | 0.03 | < 0.1
| Mean | [10 empty] | 0.015 |  | 0.03 | 0.02 |  |  | 0.01 | 0.03
| Median | [10 empty] | 0.015 |  | 0.03 | 0.02 |  |  | 0.01 | 0.03
| Std Dev. | [10 empty] | 0.0070711 |  | 0.0063246
| Rel.Std.Dev. | [10 empty] | 47.14% |  | 21.08%
PDM3
| Z-Score (Absolute) |  | 2.27 | 366.39 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 2.27 | 0.18 | 2.27 | 3 | 0.31 | 2.27 | 20.66 | 0.06 | 0.06 | 5.95
| Analytical results for Tb in OREAS 313 (Certified Value 0.31 ppm) | [25 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.12 | 2.13 | 2.21 | 2.22 | 2.23 | [16 empty] | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.25g | 0.2g | 01g | [16 empty] | 3
| 1 | 1 | 0.2953 | < 5 | 0.33323 | 0.3 | 0.29 | 0.29 | 0.34 | 0.36 | [16 empty] | 1
| 1 | 2 | 0.2881 | < 5 | 0.31338 | 0.3 | 0.28 | 0.28 | 0.35 | 0.36 | [16 empty] | 12
| 1 | 3 | 0.281 | < 5 | 0.33611 | 0.3 | 0.3 | 0.29 | 0.35 | 0.35 | [16 empty] | 16
| 1 | 4 | 0.2957 | < 5 | 0.32257 | 0.3 | 0.33 | 0.28 | 0.33 | 0.34 | [16 empty] | 0.31448
| 1 | 5 | 0.2689 | < 5 | 0.34707 | 0.3 | 0.27 | 0.3 | 0.34 | 0.35 | [16 empty] | 104
| 1 | 6 | 0.2777 | < 5 | 0.33212 | 0.3 | 0.26 | 0.29 | 0.34 | 0.36
| Mean |  | 0.28445 |  | 0.33075 | 0.3 | 0.28833 | 0.28833 | 0.34167 | 0.35333
| Median |  | 0.28455 |  | 0.33268 | 0.3 | 0.285 | 0.29 | 0.34 | 0.355
| Std Dev. |  | 0.010553 |  | 0.011591 | 0 | 0.024833 | 0.0075277 | 0.0075277 | 0.008165
| Rel.Std.Dev. |  | 3.71% |  | 3.504% | 0% | 8.613% | 2.611% | 2.203% | 2.311%
| PDM3 |  | -9.548% |  | 5.173% | -4.604% | -8.313% | -8.313% | 8.646% | 12.36%
| Z-Score (Absolute) |  | 0.74 | 34.49 | 0 | NA | 0.67 | 0.67 | 0.17 | 0.36
Results from laboratory 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 313 (Certified Value 0.069 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.07 | 0.06 | < 5 | 0.07 | < 0.2 | 0.07 | 0.09 | 0.07 | 0.08 | 0.06 | 0.05 | 0.06 | 0.063611 | 0.06 | 0.11 | 0.08 | < 0.2 | 3.57 | 0.24 |  | 0.1 | < 0.1 |  |  | 1
| 1 | 2 | 0.06 | 0.06 | < 5 | 0.07 | < 0.2 | 0.08 | 0.09 | 0.07 | 0.07 | 0.07 | 0.05 | 0.07 | 0.072145 | < 0.01 | 0.08 | 0.08 | < 0.2 | 3.57 | 0.2 | 0.08 | 0.1 | < 0.1 |  |  | 13
| 1 | 3 | 0.05 | 0.06 | < 5 | 0.07 | < 0.2 | 0.07 | 0.05 | 0.08 | 0.07 | 0.07 | 0.05 | 0.08 | 0.096536 | 0.1 | 0.12 | < 0.03 | < 0.2 | 3.47 | 0.24 |  | 0.08 | < 0.1 |  |  | 16
| 1 | 4 | 0.06 | 0.06 | < 5 | 0.06 | < 0.2 | 0.07 | 0.07 | 0.09 | 0.07 | 0.07 | 0.06 | 0.07 | 0.075035 | < 0.01 | 0.04 | 0.08 | < 0.2 | 3.51 | 0.26 |  | 0.08 | < 0.1 |  |  | 0.0685
| 1 | 5 | 0.05 | 0.05 | < 5 | 0.07 | < 0.2 | 0.06 | 0.05 | 0.08 | 0.06 | 0.06 | 0.07 | 0.06 | 0.090341 | < 0.01 | 0.05 | 0.07 | < 0.2 | 3.76 | 0.35 | 0.06 | 0.08 | < 0.1 |  |  | 105
| 1 | 6 | 0.06 | 0.07 | < 5 | 0.07 | < 0.2 | 0.07 | 0.07 | 0.08 | 0.06 | 0.06 | 0.05 | 0.05 | 0.08935 | 0.02 | 0.04 | 0.06 | < 0.2 | 3.51 | 0.26 |  | 0.09 | < 0.1
| Mean |  | 0.058333 | 0.06 |  | 0.068333 |  | 0.07 | 0.07 | 0.078333 | 0.068333 | 0.065 | 0.055 | 0.065 | 0.08117 | 0.06 | 0.073333 | 0.074 |  | 3.565 | 0.25833 | 0.07 | 0.088333
| Median |  | 0.06 | 0.06 |  | 0.07 |  | 0.07 | 0.07 | 0.08 | 0.07 | 0.065 | 0.05 | 0.065 | 0.082193 | 0.06 | 0.065 | 0.08 |  | 3.54 | 0.25 | 0.07 | 0.085
| Std Dev. |  | 0.0075277 | 0.0063246 |  | 0.0040825 |  | 0.0063246 | 0.017889 | 0.0075277 | 0.0075277 | 0.0054772 | 0.0083666 | 0.010488 | 0.012763 | 0.04 | 0.03559 | 0.0089443 |  | 0.1031 | 0.049967 | 0.014142 | 0.0098319
| Rel.Std.Dev. |  | 12.9% | 10.54% |  | 5.974% |  | 9.035% | 25.56% | 9.61% | 11.02% | 8.427% | 15.21% | 16.14% | 15.72% | 66.67% | 48.53% | 12.09% |  | 2.892% | 19.34% | 20.2% | 11.13%
| PDM3 |  | -14.84% | -12.41% |  | -0.2436% |  | 2.189% | 2.189% | 14.35% | -0.2436% | -5.11% | -19.71% | -5.11% | 18.5% | -12.41% | 7.056% | 8.029% |  | 5104% | 277.1% | 2.189% | 28.95%
| Z-Score (Absolute) |  | 0.74 | 0.64 | 154.8 | 0.11 | 1.91 | 0 | 0 | 0.53 | 0.11 | 0.32 | 0.96 | 0.32 | 0.71 | 2.39 | 0.21 | 0.37 | 1.91 | 222.65 | 12 | 0 | 1.17 | 1.27
| Analytical results for Th in OREAS 313 (Certified Value 6.36 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 7.57 | 6.9579 | < 5 | 6.227 | 6.7378 | 6 | 5.7 | 6.4 | 7 | 6.6 | 5.8 | 6 | 7.3167 | 4.3 | 6 | 7 | 4.3 | 5.4 | 8.23 | 6.6 | 4.7 |  |  |  | 1
| 1 | 2 | 7.64 | 6.9468 | < 5 | 6.407 | 6.8701 | 6 | 5.8 | 6.2 | 6.9 | 6.9 | 5.6 | 6.49 | 7.1179 | 4.5 | 6 | 6.8 | 4.5 | 5.3 | 8.18 | 6.5 | 4.6 |  |  |  | 14
| 1 | 3 | 7.63 | 6.9475 | < 5 | 6.168 | 6.7639 | 6.4 | 5.8 | 6.5 | 6.7 | 6.8 | 5.5 | 6.32 | 7.2451 | 4.6 | 6 | 5.7 | 4.4 | 5.3 | 8.23 | 6.3 | 4.6 |  |  |  | 16
| 1 | 4 | 7.6 | 6.8846 | < 5 | 6.412 | 6.5569 | 6 | 5.7 | 6.3 | 6.9 | 6.7 | 5.5 | 6.94 | 7.2994 | 4.6 | 6.4 | 7.1 | 4.4 | 5.2 | 8.03 | 6.4 | 4.5 |  |  |  | 6.3637
| 1 | 5 | 7.55 | 6.981 | < 5 | 6.485 | 7.2104 | 6.2 | 5.6 | 6.4 | 6.8 | 6.9 | 5.8 | 6.07 | 7.2479 | 4.5 | 6 | 6.7 | 4.4 | 5.4 | 8.02 | 6.6 | 4.6 |  |  |  | 106
| 1 | 6 | 7.63 | 6.9263 | < 5 | 6.726 | 7.2483 | 6 | 5.8 | 6.5 | 7 | 6.7 | 6.3 | 5.79 | 7.119 | 4.8 | 5.7 | 6.9 | 4.4 | 5.2 | 8.26 | 6.8 | 4.8
| Mean |  | 7.6033 | 6.9407 |  | 6.4042 | 6.8979 | 6.1 | 5.7333 | 6.3833 | 6.8833 | 6.7667 | 5.75 | 6.2683 | 7.2243 | 4.55 | 6.0167 | 6.7 | 4.4 | 5.3 | 8.1583 | 6.5333 | 4.6333
| Median |  | 7.615 | 6.9471 |  | 6.4095 | 6.817 | 6 | 5.75 | 6.4 | 6.9 | 6.75 | 5.7 | 6.195 | 7.2465 | 4.55 | 6 | 6.85 | 4.4 | 5.3 | 8.205 | 6.55 | 4.6
| Std Dev. |  | 0.036697 | 0.032735 |  | 0.19867 | 0.27605 | 0.16733 | 0.08165 | 0.1169 | 0.1169 | 0.12111 | 0.30166 | 0.41063 | 0.08669 | 0.16432 | 0.22286 | 0.5099 | 0.063246 | 0.089443 | 0.10647 | 0.17512 | 0.10328
| Rel.Std.Dev. |  | 0.4826% | 0.4716% |  | 3.102% | 4.002% | 2.743% | 1.424% | 1.831% | 1.698% | 1.79% | 5.246% | 6.551% | 1.2% | 3.611% | 3.704% | 7.61% | 1.437% | 1.688% | 1.305% | 2.68% | 2.229%
| PDM3 |  | 19.48% | 9.067% |  | 0.6358% | 8.394% | -4.144% | -9.906% | 0.3084% | 8.165% | 6.332% | -9.644% | -1.499% | 13.52% | -28.5% | -5.453% | 5.285% | -30.86% | -16.72% | 28.2% | 2.666% | -27.19%
| Z-Score (Absolute) |  | 1.48 | 0.67 | 4.7 | 0.03 | 0.62 | 0.34 | 0.79 | 0 | 0.6 | 0.46 | 0.77 | 0.14 | 1.02 | 2.22 | 0.44 | 0.38 | 2.4 | 1.31 | 2.15 | 0.18 | 2.12
| Analytical results for Ti in OREAS 313 (Certified Value 0.006 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.0089 | 0.005194 | 0.01 | 0.0046 | 0.0074662 | 0.005 | 0.005 | 0.006 | 0.005 | 0.005 | < 0.01 | 0.006 | 0.0095667 | < 0.01 | 0.006 | 0.004 | < 0.01 | < 0.01 | < 0.01 | 0.005 | 0.006 |  |  |  | 1
| 1 | 2 | 0.009 | 0.005292 | 0.01 | 0.005 | 0.0072376 | 0.005 | 0.005 | 0.006 | 0.005 | 0.006 | < 0.01 | 0.006 | 0.0094905 | < 0.01 | 0.005 | 0.005 | < 0.01 | < 0.01 | < 0.01 | 0.006 | 0.006 |  |  |  | 37
| 1 | 3 | 0.009 | 0.004998 | 0.01 | 0.005 | 0.0074427 | 0.005 | 0.006 | 0.006 | 0.005 | 0.006 | < 0.01 | 0.006 | 0.0099285 | < 0.01 | 0.005 | 0.004 | < 0.01 | < 0.01 | < 0.01 | 0.005 | 0.006 |  |  |  | 16
| 1 | 4 | 0.0091 | 0.005415 | 0.01 | 0.0049 | 0.0071409 | 0.006 | 0.006 | 0.006 | 0.005 | 0.006 | < 0.01 | 0.006 | 0.010104 | < 0.01 | 0.006 | 0.005 | < 0.01 | < 0.01 | < 0.01 | 0.005 | 0.006 |  |  |  | 0.0055248
| 1 | 5 | 0.0089 | 0.004802 | 0.01 | 0.005 | 0.0074422 | 0.005 | 0.005 | 0.006 | 0.005 | 0.006 | < 0.01 | 0.006 | 0.01027 | < 0.01 | 0.004 | 0.004 | < 0.01 | < 0.01 | < 0.01 | 0.006 | 0.006 |  |  |  | 107
| 1 | 6 | 0.0088 | 0.005096 | 0.01 | 0.0053 | 0.0076083 | 0.005 | 0.005 | 0.006 | 0.005 | 0.006 | < 0.01 | 0.006 | 0.00999 | < 0.01 | 0.005 | 0.004 | < 0.01 | < 0.01 | < 0.01 | 0.006 | 0.006
| Mean |  | 0.00895 | 0.0051328 | 0.01 | 0.0049667 | 0.0073897 | 0.0051667 | 0.0053333 | 0.006 | 0.005 | 0.0058333 |  | 0.006 | 0.0098915 |  | 0.0051667 | 0.0043333 |  |  |  | 0.0055 | 0.006
| Median |  | 0.00895 | 0.005145 | 0.01 | 0.005 | 0.0074424 | 0.005 | 0.005 | 0.006 | 0.005 | 0.006 |  | 0.006 | 0.0099593 |  | 0.005 | 0.004 |  |  |  | 0.0055 | 0.006
| Std Dev. |  | 0.00010488 | 0.00021803 | 0 | 0.00022509 | 0.00016986 | 0.00040825 | 0.0005164 | 9.5015e-19 | 0 | 0.00040825 |  | 9.5015e-19 | 0.00030509 |  | 0.00075277 | 0.0005164 |  |  |  | 0.00054772 | 9.5015e-19
| Rel.Std.Dev. |  | 1.172% | 4.248% | 0% | 4.532% | 2.299% | 7.902% | 9.682% | 1.584e-14% | 0% | 6.999% |  | 1.584e-14% | 3.084% |  | 14.57% | 11.92% |  |  |  | 9.959% | 1.584e-14%
| PDM3 |  | 62% | -7.095% | 81% | -10.1% | 33.75% | -6.482% | -3.466% | 8.601% | -9.499% | 5.584% |  | 8.601% | 79.04% |  | -6.482% | -21.57% |  |  |  | -0.449% | 8.601%
| Z-Score (Absolute) |  | 12.76 | 0.11 | 16.3 | 0.67 | 7.49 | 0 | 0.56 | 2.81 | 0.56 | 2.25 | 0.56 | 2.81 | 15.93 | 0.56 | 0 | 2.81 | 0.56 | 0.56 | 0.56 | 1.12 | 2.81
| Analytical results for Tl in OREAS 313 (Certified Value 10.9 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 11.3 | 11.23 | < 5 | 11.07 | 14.2043 | 10.6 | 10.7 | 11.65 | 10.65 | 10.65 | 10.1 | 11.06 | 11.4462 | 10.4 | 10.71 | 9.042 | 11.3 | 10 | 8.23 | 10.98 | 12.6 | 11 | 7.79 |  | 1
| 1 | 2 | 11.2 | 12.05 | < 5 | 11.3 | 13.506 | 10.75 | 11 | 11.15 | 10.75 | 10.75 | 9.59 | 11.62 | 11.3713 | 10.6 | 11.07 | 8.873 | 11.4 | 10 | 8.16 | 11.16 | 12.34 | 10.65 | 7.59 |  | 15
| 1 | 3 | 11.2 | 11.151 | < 5 | 10.87 | 14.4832 | 11 | 10.95 | 11.7 | 10.7 | 10.65 | 9.74 | 11.6 | 11.35 | 10.1 | 10.87 | 9.16 | 9.3 | 10 | 8.03 | 10.77 | 12.46 | 10.38 | 7.67 |  | 16
| 1 | 4 | 11.1 | 11.32 | < 5 | 11.19 | 13.9213 | 10.5 | 10.5 | 11.8 | 10.3 | 10.65 | 9.51 | 11.98 | 11.4554 | 10.6 | 11.27 | 9.023 | 11.9 | < 10 | 8.13 | 10.6 | 12.16 | 10.97 | 7.55 |  | 10.9262
| 1 | 5 | 11.1 | 11.36 | < 5 | 11.43 | 14.5731 | 10.75 | 10.55 | 11.8 | 10.45 | 10.6 | 9.91 | 11.35 | 11.4151 | 11.3 | 10.92 | 9.353 | 11 | 10 | 8.3 | 10.61 | 12.09 | 10.72 | 7.59 |  | 108
| 1 | 6 | 11.2 | 12.3 | < 5 | 11.97 | 14.6937 | 10.5 | 10.65 | 11.05 | 10.8 | 10.7 | 9.18 | 11.18 | 11.3713 | 10.7 | 10.56 | 9.042 | 11.2 | < 10 | 8.05 | 10.69 | 12.64 | 10.96 | 7.67
| Mean |  | 11.1833 | 11.5685 |  | 11.305 | 14.2302 | 10.6833 | 10.725 | 11.525 | 10.6083 | 10.6667 | 9.6717 | 11.465 | 11.4016 | 10.6167 | 10.9 | 9.0822 | 11.0167 | 10 | 8.15 | 10.8017 | 12.3817 | 10.78 | 7.6433
| Median |  | 11.2 | 11.34 |  | 11.245 | 14.3437 | 10.675 | 10.675 | 11.675 | 10.675 | 10.65 | 9.665 | 11.475 | 11.3932 | 10.6 | 10.895 | 9.042 | 11.25 | 10 | 8.145 | 10.73 | 12.4 | 10.84 | 7.63
| Std Dev. |  | 0.075277 | 0.48188 |  | 0.3783 | 0.451 | 0.19149 | 0.20676 | 0.33578 | 0.19343 | 0.05164 | 0.32233 | 0.33644 | 0.043728 | 0.39707 | 0.25251 | 0.16117 | 0.89312 | 0 | 0.10373 | 0.22409 | 0.22649 | 0.24389 | 0.08641
| Rel.Std.Dev. |  | 0.6731% | 4.165% |  | 3.346% | 3.169% | 1.792% | 1.928% | 2.914% | 1.823% | 0.4841% | 3.333% | 2.934% | 0.3835% | 3.74% | 2.317% | 1.775% | 8.107% | 0% | 1.273% | 2.075% | 1.829% | 2.262% | 1.131%
| PDM3 |  | 2.353% | 5.879% |  | 3.467% | 30.24% | -2.223% | -1.841% | 5.481% | -2.909% | -2.375% | -11.48% | 4.931% | 4.351% | -2.833% | -0.2397% | -16.88% | 0.8281% | -8.477% | -25.41% | -1.14% | 13.32% | -1.338% | -30.05%
| Z-Score (Absolute) |  | 0.47 | 0.93 | 9.94 | 0.62 | 4.12 | 0.13 | 0.08 | 0.88 | 0.22 | 0.15 | 1.34 | 0.81 | 0.73 | 0.21 | 0.13 | 2.05 | 0.27 | NA | 3.17 | 0.01 | 1.91 | 0.01 | 3.77
Results from laboratory 2.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Tm in OREAS 313 (Indicative Value 0.089 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | [21 empty] | 1
| No | No | 2.06 | 2.12 | 2.23 | [21 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [21 empty] | 3
|  |  | 0.2g | 0.5g | 01g | [21 empty] | 3
| 1 | 1 | 0.078895 | < 0.1 | 0.1 | [21 empty] | 1
| 1 | 2 | 0.072517 | < 0.1 | 0.1 | [21 empty] | 8
| 1 | 3 | 0.080053 | < 0.1 | 0.1 | [21 empty] | 16
| 1 | 4 | 0.068945 | < 0.1 | 0.1 | [21 empty] | 0.08867
| 1 | 5 | 0.082125 | < 0.1 | 0.1 | [21 empty] | 14
| 1 | 6 | 0.081505 | < 0.1 | 0.1
| Mean |  | 0.07734 |  | 0.1
| Median |  | 0.079474 |  | 0.1
| Std Dev. |  | 0.005362 |  | 1.5202e-17
| Rel.Std.Dev. |  | 6.933% |  | 1.52e-14%
| PDM3 |  | -12.78% |  | 12.78%
| Z-Score (Absolute) |  | 0 | 0.81 | 0.67
| Analytical results for U in OREAS 313 (Certified Value 1.81 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 2 | 1.5279 | 1.74 | < 5 | 1.901 | 1.7811 | 1.66 | 1.8 | 1.82 | 1.66 | 1.78 | 1.8 | 1.73 | 2.1312 | 1.69 | 1.84 | 1.57 | 2 | 1.35 | 1.59 | 1.79 | 2.05 | 1.96 | 1.3 | 1
| 1 | 2 | 1.96 | 1.5048 | 1.87 | < 5 | 1.948 | 1.7799 | 1.74 | 1.84 | 1.78 | 1.72 | 1.81 | 1.8 | 1.83 | 2.0456 | 1.75 | 1.89 | 1.53 | 1.9 | 1.4 | 1.54 | 1.88 | 2.08 | 1.88 | 1.3 | 17
| 1 | 3 | 1.99 | 1.4706 | 1.7876 | < 5 | 1.9 | 1.9269 | 1.78 | 1.88 | 1.8 | 1.67 | 1.81 | 1.8 | 1.89 | 2.079 | 1.71 | 1.9 | 1.59 | 1.6 | 1.42 | 1.54 | 1.72 | 2.07 | 1.82 | 1.3 | 16
| 1 | 4 | 1.98 | 1.516 | 1.83 | < 5 | 1.964 | 1.7534 | 1.68 | 1.78 | 1.88 | 1.7 | 1.79 | 1.7 | 1.91 | 2.1774 | 1.73 | 2 | 1.6 | 2 | 1.39 | 1.49 | 1.82 | 2 | 1.93 | 1.3 | 1.8139
| 1 | 5 | 1.97 | 1.4452 | 1.78 | < 5 | 1.969 | 1.9612 | 1.73 | 1.82 | 1.89 | 1.7 | 1.8 | 1.8 | 1.71 | 2.1563 | 1.84 | 1.87 | 1.6 | 1.9 | 1.41 | 1.57 | 1.86 | 2.03 | 1.91 | 1.3 | 109
| 1 | 6 | 1.97 | 1.5135 | 1.9 | < 5 | 2.101 | 1.8989 | 1.68 | 1.82 | 1.93 | 1.68 | 1.79 | 1.8 | 1.66 | 2.084 | 1.77 | 1.77 | 1.56 | 1.9 | 1.38 | 1.5 | 1.85 | 2.08 | 1.94 | 1.3
| Mean |  | 1.9783 | 1.4963 | 1.8179 |  | 1.9638 | 1.8502 | 1.7117 | 1.8233 | 1.85 | 1.6883 | 1.7967 | 1.7833 | 1.7883 | 2.1123 | 1.7483 | 1.8783 | 1.575 | 1.8833 | 1.3917 | 1.5383 | 1.82 | 2.0517 | 1.9067 | 1.3
| Median |  | 1.975 | 1.5091 | 1.8088 |  | 1.956 | 1.84 | 1.705 | 1.82 | 1.85 | 1.69 | 1.795 | 1.8 | 1.78 | 2.1076 | 1.74 | 1.88 | 1.58 | 1.9 | 1.395 | 1.54 | 1.835 | 2.06 | 1.92 | 1.3
| Std Dev. |  | 0.01472 | 0.031706 | 0.06004 |  | 0.073641 | 0.089073 | 0.045789 | 0.034448 | 0.058652 | 0.022286 | 0.012111 | 0.040825 | 0.10284 | 0.050766 | 0.053072 | 0.075741 | 0.027386 | 0.1472 | 0.024833 | 0.038687 | 0.05831 | 0.031885 | 0.050465 | 0
| Rel.Std.Dev. |  | 0.744% | 2.119% | 3.303% |  | 3.75% | 4.814% | 2.675% | 1.889% | 3.17% | 1.32% | 0.6741% | 2.289% | 5.751% | 2.403% | 3.036% | 4.032% | 1.739% | 7.816% | 1.784% | 2.515% | 3.204% | 1.554% | 2.647% | 0%
| PDM3 |  | 9.066% | -17.51% | 0.2235% |  | 8.267% | 2.003% | -5.635% | 0.5212% | 1.991% | -6.921% | -0.949% | -1.684% | -1.408% | 16.45% | -3.614% | 3.553% | -13.17% | 3.829% | -23.28% | -15.19% | 0.3374% | 13.11% | 5.115% | -28.33%
| Z-Score (Absolute) |  | 1.1 | 2.23 | 0.01 | 4.71 | 1 | 0.22 | 0.74 | 0.03 | 0.21 | 0.9 | 0.15 | 0.25 | 0.21 | 2.03 | 0.49 | 0.41 | 1.69 | 0.45 | 2.96 | 1.94 | 0.01 | 1.61 | 0.61 | 3.59
| Analytical results for V in OREAS 313 (Certified Value 17.8 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 24 | 28.933 | 16 | 16.2 | 17 | 20.02 | 17 | 16 | 19 | 16 | 17 | 19 | 18 | 19.3679 | 15 | 20 | 18 | 18 | 17 | 18.5 | 21 | 17 |  |  | 1
| 1 | 2 | 25 | 25.4365 | 15 | 17 | 17.6 | 19.6207 | 18 | 17 | 18 | 16 | 18 | 20 | 17 | 20.8666 | 16 | 20 | 18 | 18 | 17 | 18.8 | 22 | 16 |  |  | 18
| 1 | 3 | 25 | 25.6885 | 14 | 15.3 | 17.2 | 19.8394 | 17 | 17 | 18 | 16 | 17 | 20 | 17 | 20.4242 | 15 | 19 | 16 | 18 | 17 | 18.3 | 21 | 16 |  |  | 16
| 1 | 4 | 25 | 25.93 | 15 | 17.6 | 17.3 | 19.6252 | 18 | 17 | 19 | 16 | 17 | 18 | 18 | 20.1813 | 15 | 21 | 19 | 18 | 17 | 18.2 | 21 | 16 |  |  | 17.7802
| 1 | 5 | 24 | 27.0745 | 15 | 16.7 | 17.5 | 20.334 | 17 | 17 | 19 | 16 | 18 | 20 | 18 | 19.3128 | 16 | 18 | 17 | 18 | 18 | 18.2 | 21 | 16 |  |  | 110
| 1 | 6 | 24 | 23.4625 | 15 | 16.7 | 18.5 | 20.6151 | 17 | 17 | 19 | 16 | 17 | 21 | 17 | 18.5991 | 16 | 18 | 18 | 17 | 17 | 18.4 | 22 | 17
| Mean |  | 24.5 | 26.0875 | 15 | 16.5833 | 17.5167 | 20.0091 | 17.3333 | 16.8333 | 18.6667 | 16 | 17.3333 | 19.6667 | 17.5 | 19.792 | 15.5 | 19.3333 | 17.6667 | 17.8333 | 17.1667 | 18.4 | 21.3333 | 16.3333
| Median |  | 24.5 | 25.8093 | 15 | 16.7 | 17.4 | 19.9297 | 17 | 17 | 19 | 16 | 17 | 20 | 17.5 | 19.7746 | 15.5 | 19.5 | 18 | 18 | 17 | 18.35 | 21 | 16
| Std Dev. |  | 0.54772 | 1.8204 | 0.63246 | 0.77825 | 0.52694 | 0.39998 | 0.5164 | 0.40825 | 0.5164 | 0 | 0.5164 | 1.0328 | 0.54772 | 0.84125 | 0.54772 | 1.2111 | 1.0328 | 0.40825 | 0.40825 | 0.22804 | 0.5164 | 0.5164
| Rel.Std.Dev. |  | 2.236% | 6.978% | 4.216% | 4.693% | 3.008% | 1.999% | 2.979% | 2.425% | 2.766% | 0% | 2.979% | 5.252% | 3.13% | 4.25% | 3.534% | 6.264% | 5.846% | 2.289% | 2.378% | 1.239% | 2.421% | 3.162%
| PDM3 |  | 37.79% | 46.72% | -15.64% | -6.732% | -1.482% | 12.54% | -2.513% | -5.325% | 4.986% | -10.01% | -2.513% | 10.61% | -1.576% | 11.31% | -12.82% | 8.735% | -0.6386% | 0.2987% | -3.451% | 3.486% | 19.98% | -8.138%
| Z-Score (Absolute) |  | 3.99 | 4.91 | 1.5 | 0.58 | 0.04 | 1.4 | 0.15 | 0.44 | 0.62 | 0.92 | 0.15 | 1.2 | 0.05 | 1.27 | 1.21 | 1.01 | 0.04 | 0.14 | 0.25 | 0.47 | 2.16 | 0.73
| Analytical results for W in OREAS 313 (Certified Value 0.14 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 0.2 | 0.23 | < 5 | 0.11 | 0.2467 | 0.13 | 0.13 | 0.12 | 0.12 | 0.14 | 0.2 | 0.17 | 0.31855 | 0.87 | 0.13 | 0.1 | 0.38 | 0.4 | 0.17 | 0.16 | 0.18 | < 0.1 |  |  | 1
| 1 | 2 | 0.2 | 0.24 | < 5 | 0.13 | 0.23196 | 0.14 | 0.15 | 0.12 | 0.12 | 0.14 | 0.2 | 0.14 | 0.33155 | 0.38 | 0.13 | 0.1 | 0.38 | 0.4 | 0.19 | 0.17 | 0.17 | < 0.1 |  |  | 19
| 1 | 3 | 0.2 | 0.23 | < 5 | 0.13 | 0.2457 | 0.14 | 0.13 | 0.12 | 0.13 | 0.14 | 0.2 | 0.14 | 0.29088 | 0.6 | 0.13 | 0.6 | 0.39 | 0.4 | 0.15 | 0.16 | 0.16 | < 0.1 |  |  | 16
| 1 | 4 | 0.2 | 0.22 | < 5 | 0.13 | 0.23628 | 0.13 | 0.14 | 0.12 | 0.13 | 0.14 | 0.2 | 0.14 | 0.27832 | 0.45 | 0.13 | 0.1 | 0.37 | 0.3 | 0.16 | 0.15 | 0.16 | < 0.1 |  |  | 0.14203
| 1 | 5 | 0.2 | 0.23 | < 5 | 0.15 | 0.25173 | 0.14 | 0.12 | 0.12 | 0.13 | 0.14 | 0.2 | 0.14 | 0.32681 | 0.32 | 0.15 | 0.1 | 0.39 | 0.4 | 0.17 | 0.15 | 0.17 | < 0.1 |  |  | 111
| 1 | 6 | 0.2 | 0.23 | < 5 | 0.14 | 0.23019 | 0.14 | 0.13 | 0.12 | 0.13 | 0.14 | 0.2 | 0.14 | 0.28666 | 0.5 | 0.39 | 0.1 | 0.33 | 0.3 | 0.19 | 0.16 | 0.18 | 0.1
| Mean |  | 0.2 | 0.23 |  | 0.13167 | 0.24043 | 0.13667 | 0.13333 | 0.12 | 0.12667 | 0.14 | 0.2 | 0.145 | 0.30546 | 0.52 | 0.17667 | 0.18333 | 0.37333 | 0.36667 | 0.17167 | 0.15833 | 0.17 | 0.1
| Median |  | 0.2 | 0.23 |  | 0.13 | 0.24099 | 0.14 | 0.13 | 0.12 | 0.13 | 0.14 | 0.2 | 0.14 | 0.30472 | 0.475 | 0.13 | 0.1 | 0.38 | 0.4 | 0.17 | 0.16 | 0.17 | 0.1
| Std Dev. |  | 3.0405e-17 | 0.0063246 |  | 0.013292 | 0.0088162 | 0.005164 | 0.010328 | 0 | 0.005164 | 0 | 3.0405e-17 | 0.012247 | 0.02285 | 0.19688 | 0.10482 | 0.20412 | 0.022509 | 0.05164 | 0.016021 | 0.0075277 | 0.0089443
| Rel.Std.Dev. |  | 1.52e-14% | 2.75% |  | 10.09% | 3.667% | 3.779% | 7.746% | 0% | 4.077% | 0% | 1.52e-14% | 8.447% | 7.481% | 37.86% | 59.33% | 111.3% | 6.029% | 14.08% | 9.333% | 4.754% | 5.261%
| PDM3 |  | 40.82% | 61.94% |  | -7.297% | 69.28% | -3.776% | -6.123% | -15.51% | -10.82% | -1.429% | 40.82% | 2.091% | 115.1% | 266.1% | 24.39% | 29.08% | 162.9% | 158.2% | 20.87% | 11.48% | 19.69% | -29.59%
| Z-Score (Absolute) |  | NA | 1.06 | 40.99 | 0.67 | 1.24 | 0.59 | 0.64 | 0.88 | 0.76 | 0.53 | NA | 0.44 | 2.38 | 6.16 | 0.12 | NA | 3.58 | NA | 0.03 | 0.21 | 0 | NA
Results from laboratories 2.01, 2.12, 2.19, 2.21 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 313 (Certified Value 5.63 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 5.84 | 5.86 | 5.8 | 5.41 | 6.1121 | 5.69 | 5.18 | 5.65 | 5.4 | 5.77 | 4.78 | 5.78 | 6.1126 | 5.08 | 3.47 | 5.45 | 5.3 | 5.98 | 5.92 | 3.9 |  |  |  |  | 1
| 1 | 2 | 5.81 | 5.9364 | 5.5 | 5.54 | 6.0103 | 5.67 | 4.9 | 5.71 | 5.4 | 5.85 | 4.81 | 5.64 | 6.2438 | 4.99 | 3.57 | 5.53 | 5.6 | 6.14 | 5.85 | 3.8 |  |  |  |  | 20
| 1 | 3 | 5.8 | 5.9683 | 6.3 | 5.44 | 6.0782 | 5.68 | 4.83 | 5.7 | 5.31 | 5.7 | 5 | 5.83 | 6.0408 | 5.19 | 3.63 | 5.44 | 5.3 | 5.98 | 5.63 | 3.9 |  |  |  |  | 16
| 1 | 4 | 5.91 | 5.8471 | 6.3 | 5.34 | 6.0044 | 5.59 | 4.86 | 6.19 | 5.42 | 5.75 | 4.44 | 5.93 | 6.2882 | 5.07 | 3.55 | 5.18 | 5.3 | 6.01 | 5.7 | 3.9 |  |  |  |  | 5.6309
| 1 | 5 | 5.82 | 5.8727 | 5.7 | 5.44 | 6.148 | 5.63 | 4.9 | 6.29 | 5.23 | 5.87 | 4.89 | 5.79 | 6.1443 | 5.38 | 3.57 | 5.41 | 5.3 | 5.74 | 5.94 | 3.9 |  |  |  |  | 112
| 1 | 6 | 5.93 | 5.7163 | 6.3 | 5.71 | 6.2549 | 5.56 | 4.89 | 5.59 | 5.59 | 5.82 | 4.81 | 5.51 | 6.1645 | 5.35 | 3.45 | 5.35 | 5.4 | 6.1 | 6.17 | 4
| Mean |  | 5.8517 | 5.8668 | 5.9833 | 5.48 | 6.1013 | 5.6367 | 4.9267 | 5.855 | 5.3917 | 5.7933 | 4.7883 | 5.7467 | 6.1657 | 5.1767 | 3.54 | 5.3933 | 5.3667 | 5.9917 | 5.8683 | 3.9
| Median |  | 5.83 | 5.8664 | 6.05 | 5.44 | 6.0951 | 5.65 | 4.895 | 5.705 | 5.4 | 5.795 | 4.81 | 5.785 | 6.1544 | 5.135 | 3.56 | 5.425 | 5.3 | 5.995 | 5.885 | 3.9
| Std Dev. |  | 0.054924 | 0.08748 | 0.36009 | 0.12977 | 0.093899 | 0.052789 | 0.12707 | 0.3029 | 0.1209 | 0.064704 | 0.18841 | 0.14895 | 0.089435 | 0.15946 | 0.067823 | 0.11978 | 0.12111 | 0.13977 | 0.19198 | 0.063246
| Rel.Std.Dev. |  | 0.9386% | 1.491% | 6.018% | 2.368% | 1.539% | 0.9365% | 2.579% | 5.173% | 2.242% | 1.117% | 3.935% | 2.592% | 1.451% | 3.08% | 1.916% | 2.221% | 2.257% | 2.333% | 3.271% | 1.622%
| PDM3 |  | 3.921% | 4.19% | 6.259% | -2.68% | 8.354% | 0.1025% | -12.51% | 3.98% | -4.249% | 2.885% | -14.96% | 2.056% | 9.497% | -8.067% | -37.13% | -4.219% | -4.693% | 6.407% | 4.217% | -30.74%
| Z-Score (Absolute) |  | 0.36 | 0.39 | 0.66 | 0.48 | 0.92 | 0.12 | 1.72 | 0.37 | 0.68 | 0.23 | 2.04 | 0.12 | 1.07 | 1.16 | 4.85 | 0.67 | 0.73 | 0.68 | 0.4 | 4.04
| Analytical results for Yb in OREAS 313 (Certified Value 0.56 ppm) | [25 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.12 | 2.13 | 2.21 | 2.22 | 2.23 | [16 empty] | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.25g | 0.2g | 01g | [16 empty] | 3
| 1 | 1 | 0.519 | < 5 | 0.58507 | 0.6 | 0.6 | 0.6 | 0.51 | 0.68 | [16 empty] | 1
| 1 | 2 | 0.532 | < 5 | 0.56788 | 0.6 | 0.5 | 0.6 | 0.55 | 0.69 | [16 empty] | 21
| 1 | 3 | 0.5202 | < 5 | 0.60038 | 0.5 | 0.6 | 0.6 | 0.52 | 0.68 | [16 empty] | 16
| 1 | 4 | 0.5225 | < 5 | 0.57759 | 0.5 | 0.6 | 0.6 | 0.5 | 0.67 | [16 empty] | 0.55648
| 1 | 5 | 0.5272 | < 5 | 0.60875 | 0.6 | 0.5 | 0.6 | 0.55 | 0.68 | [16 empty] | 113
| 1 | 6 | 0.5034 | < 5 | 0.59927 | 0.5 | 0.5 | 0.6 | 0.54 | 0.69
| Mean |  | 0.52072 |  | 0.58983 | 0.55 | 0.55 | 0.6 | 0.52833 | 0.68167
| Median |  | 0.52135 |  | 0.59217 | 0.55 | 0.55 | 0.6 | 0.53 | 0.68
| Std Dev. |  | 0.009753 |  | 0.015573 | 0.054772 | 0.054772 | 0 | 0.02137 | 0.0075277
| Rel.Std.Dev. |  | 1.873% |  | 2.64% | 9.959% | 9.959% | 0% | 4.045% | 1.104%
| PDM3 |  | -6.427% |  | 5.992% | -1.164% | -1.164% | 7.821% | -5.058% | 22.5%
| Z-Score (Absolute) |  | 0.93 | 36.32 | 0.37 | 0.37 | 0.37 | 0.57 | 0.78 | 2.1
| Analytical results for Zn in OREAS 313 (Certified Value 3.45 wt.%) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*OES | AR*OES | AR*OES | AR*AAS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*MS | AR*OES | AR*OES/MS | 2
|  |  | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 01g |  | 0.25g | 0.25g | 01g | 0.4g | 0.25g | 3
| 1 | 1 | > 2 | 3.3472 | 3.5516 | 3.56 | > 2 | 3.4486 | 3.41 | 3.43 | 3.48 | 3.46 | 3.46 | 3.68 | 3.42 | 3.5454 | 3.13 | 3.451 | 3.535 | 3.44 | 3.37 | > 1 | 3.49 | > 1 | 3.414 | > 1 | 1
| 1 | 2 | > 2 | 3.3473 | 3.5022 | 3.58 | > 2 | 3.4122 | 3.43 | 3.44 | 3.45 | 3.43 | 3.42 | 3.74 | 3.4 | 3.5256 | 3.19 | 3.41 | 3.5333 | 3.44 | 3.47 | > 1 | 3.53 | > 1 | 3.357 | > 1 | #N/A
| 1 | 3 | > 2 | 3.3288 | 3.5175 | 3.52 | > 2 | 3.4576 | 3.45 | 3.42 | 3.39 | 3.45 | 3.44 | 3.6 | 3.37 | 3.5438 | 3.24 | 3.478 | 3.6005 | 3.37 | 3.46 | > 1 | 3.53 | > 1 | 3.413 | > 1 | 16
| 1 | 4 | > 2 | 3.335 | 3.5226 | 3.47 | > 2 | 3.472 | 3.43 | 3.43 | 3.47 | 3.47 | 3.45 | 3.68 | 3.45 | 3.53 | 3.13 | 3.503 | 3.4579 | 3.43 | 3.44 | > 1 | 3.45 | > 1 | 3.395 | > 1 | 3.4491
| 1 | 5 | > 2 | 3.3588 | 3.5292 | 3.68 | > 2 | 3.4833 | 3.44 | 3.41 | 3.33 | 3.52 | 3.5 | 3.71 | 3.42 | 3.5288 | 3.11 | 3.388 | 3.5521 | 3.44 | 3.43 | > 1 | 3.48 | > 1 | 3.384 | > 1 | 114
| 1 | 6 | > 2 | 3.3655 | 3.5557 | 3.53 | > 2 | 3.4679 | 3.4 | 3.43 | 3.48 | 3.48 | 3.39 | 3.64 | 3.42 | 3.5287 | 3.14 | 3.422 | 3.4947 | 3.44 | 3.38 | > 1 | 3.53 | > 1 | 3.331 | > 1
| Mean |  |  | 3.3471 | 3.5298 | 3.5567 |  | 3.4569 | 3.4267 | 3.4267 | 3.4333 | 3.4683 | 3.4433 | 3.675 | 3.4133 | 3.5337 | 3.1567 | 3.442 | 3.5289 | 3.4267 | 3.425 |  | 3.5017 |  | 3.3823
| Median |  |  | 3.3473 | 3.5259 | 3.545 |  | 3.4627 | 3.43 | 3.43 | 3.46 | 3.465 | 3.445 | 3.68 | 3.42 | 3.5294 | 3.135 | 3.4365 | 3.5342 | 3.44 | 3.435 |  | 3.51 |  | 3.3895
| Std Dev. |  |  | 0.013815 | 0.020562 | 0.071181 |  | 0.024964 | 0.018619 | 0.010328 | 0.060882 | 0.030605 | 0.037238 | 0.049699 | 0.026583 | 0.0085723 | 0.048854 | 0.043447 | 0.048851 | 0.028048 | 0.041352 |  | 0.033714 |  | 0.032813
| Rel.Std.Dev. |  |  | 0.4127% | 0.5825% | 2.001% |  | 0.7221% | 0.5434% | 0.3014% | 1.773% | 0.8824% | 1.081% | 1.352% | 0.7788% | 0.2426% | 1.548% | 1.262% | 1.384% | 0.8185% | 1.207% |  | 0.9628% |  | 0.9701%
| PDM3 |  |  | -2.958% | 2.339% | 3.118% |  | 0.2266% | -0.6507% | -0.6507% | -0.4574% | 0.5573% | -0.1675% | 6.549% | -1.037% | 2.453% | -8.479% | -0.2062% | 2.314% | -0.6507% | -0.699% |  | 1.524% |  | -1.936%
| Z-Score (Absolute) |  | NA | 2.23 | 2.06 | 2.7 | NA | 0.35 | 0.36 | 0.36 | 0.2 | 0.62 | 0.03 | 5.48 | 0.67 | 2.16 | 6.71 | 0 | 2.04 | 0.36 | 0.4 | NA | 1.4 | NA | 1.4 | NA
| Analytical results for Zr in OREAS 313 (Certified Value 11.8 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | [5 empty] | 2
| 1 | 1 | 12.4 | 11.9 | 13.06 | 12.0087 | 12.3 | 10.8 | 11.9 | 12 | 12.2 | 11.6 | 11.5 | 11.7124 | 14.8 | 13 | 10.3 | 10.1 | 11.7 | 14.4 | 10.2 | [5 empty] | 1
| 1 | 2 | 12.7 | 11.2 | 13.03 | 11.7337 | 12.4 | 11 | 12.6 | 12.2 | 12.4 | 12 | 11 | 11.1784 | 16.6 | 13 | 10.1 | 10.1 | 12.2 | 14.1 | 10.1 | [5 empty] | 22
| 1 | 3 | 12.8 | 12.5 | 12.91 | 12.0093 | 12.5 | 10.5 | 12.3 | 12.1 | 12.4 | 12.5 | 11.1 | 11.5112 | 16.1 | 13 | 10 | 9 | 11.7 | 13.5 | 10.2 | [5 empty] | 16
| 1 | 4 | 12.5 | 11.9 | 12.84 | 11.8214 | 12.2 | 10.6 | 13.1 | 12.4 | 12.4 | 11.2 | 11.4 | 11.4835 | 17.3 | 13 | 10 | 10.4 | 12 | 13.7 | 10.1 | [5 empty] | 11.8453
| 1 | 5 | 12.6 | 12.3 | 13.24 | 12.1447 | 12.3 | 10.6 | 12.9 | 12 | 12.7 | 12 | 11.3 | 11.5276 | 17.2 | 13 | 10.3 | 9.6 | 11.4 | 14 | 10.2 | [5 empty] | 115
| 1 | 6 | 12.6 | 12.5 | 13.46 | 12.2068 | 12.2 | 10.9 | 12 | 12.4 | 12.2 | 11.1 | 11.1 | 11.2146 | 17 | 13 | 10 | 9.6 | 12 | 14.9 | 10.1
| Mean |  | 12.6 | 12.05 | 13.09 | 11.9874 | 12.3167 | 10.7333 | 12.4667 | 12.1833 | 12.3833 | 11.7333 | 11.2333 | 11.438 | 16.5 | 13 | 10.1167 | 9.8 | 11.8333 | 14.1 | 10.15
| Median |  | 12.6 | 12.1 | 13.045 | 12.009 | 12.3 | 10.7 | 12.45 | 12.15 | 12.4 | 11.8 | 11.2 | 11.4974 | 16.8 | 13 | 10.05 | 9.85 | 11.85 | 14.05 | 10.15
| Std Dev. |  | 0.14142 | 0.49699 | 0.22751 | 0.18203 | 0.1169 | 0.19664 | 0.48442 | 0.18348 | 0.18348 | 0.53541 | 0.19664 | 0.20398 | 0.94234 | 0 | 0.1472 | 0.502 | 0.28752 | 0.502 | 0.054772
| Rel.Std.Dev. |  | 1.122% | 4.124% | 1.738% | 1.519% | 0.9492% | 1.832% | 3.886% | 1.506% | 1.482% | 4.563% | 1.75% | 1.783% | 5.711% | 0% | 1.455% | 5.122% | 2.43% | 3.56% | 0.5396%
| PDM3 |  | 6.371% | 1.728% | 10.51% | 1.2% | 3.979% | -9.387% | 5.246% | 2.854% | 4.542% | -0.9452% | -5.166% | -3.439% | 39.3% | 9.748% | -14.59% | -17.27% | -0.101% | 19.03% | -14.31%
| Z-Score (Absolute) |  | 0.61 | 0 | 1.15 | 0.07 | 0.29 | 1.45 | 0.46 | 0.15 | 0.37 | 0.35 | 0.9 | 0.67 | 4.9 | 1.05 | 2.13 | 2.48 | 0.24 | 2.26 | 2.09
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Ag in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 10
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 16
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 313 (Certified Value 9.53 wt.%) | [16 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 9.77 | 9.687 | 9.69 | 8.56 | 9.5234 | 9.35 | 9.46 | 9.62 | 9.58 | 9.64 | 8.3 | 9.58 | 9.693 | 9.62 |  | 1
| 1 | 2 | 9.78 | 9.619 | 9.63 | 9.2 | 9.5441 | 9.56 | 9.33 | 9.59 | 9.51 | 9.5 | 8.29 | 9.42 | 9.636 | 9.53 |  | #N/A
| 1 | 3 | 9.89 | 9.463 | 9.45 | 8.69 | 9.4603 | 9.59 | 9.36 | 9.56 | 9.62 | 9.31 | 8.22 | 9.48 | 9.655 | 9.54 |  | 16
| 1 | 4 | 9.85 | 9.357 | 9.72 | 8.83 | 9.4393 | 9.46 | 9.36 | 9.59 | 9.58 | 9.42 | 8.23 | 9.53 | 9.655 | 9.54 |  | 9.5346
| 1 | 5 | 9.87 | 9.546 | 9.73 | 8.65 | 9.4686 | 9.54 | 9.24 | 9.55 | 9.6 | 9.7 | 8.21 | 9.52 | 9.693 | 9.57 |  | 117
| 1 | 6 | 9.83 | 9.397 | 9.72 | 8.95 | 9.3857 | 9.47 | 9.62 | 9.47 | 9.57 | 9.36 | 8.22 | 9.5 | 9.674 | 9.53
| Mean |  | 9.8317 | 9.5115 | 9.6567 | 8.8133 | 9.4702 | 9.495 | 9.395 | 9.5633 | 9.5767 | 9.4883 | 8.245 | 9.505 | 9.6677 | 9.555
| Median |  | 9.84 | 9.5045 | 9.705 | 8.76 | 9.4644 | 9.505 | 9.36 | 9.575 | 9.58 | 9.46 | 8.225 | 9.51 | 9.6645 | 9.54
| Std Dev. |  | 0.048339 | 0.12875 | 0.10764 | 0.23432 | 0.057426 | 0.08735 | 0.13081 | 0.052026 | 0.037238 | 0.15549 | 0.03937 | 0.053572 | 0.02301 | 0.035071
| Rel.Std.Dev. |  | 0.4917% | 1.354% | 1.115% | 2.659% | 0.6064% | 0.92% | 1.392% | 0.544% | 0.3888% | 1.639% | 0.4775% | 0.5636% | 0.238% | 0.367%
| PDM3 |  | 3.115% | -0.2426% | 1.28% | -7.565% | -0.6757% | -0.4157% | -1.465% | 0.301% | 0.4408% | -0.4856% | -13.53% | -0.3108% | 1.395% | 0.2136%
| Z-Score (Absolute) |  | 3.53 | 0.04 | 1.62 | 7.59 | 0.42 | 0.14 | 1.24 | 0.6 | 0.75 | 0.22 | 13.79 | 0.04 | 1.74 | 0.51
| Analytical results for As in OREAS 313 (Indicative Value 105 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | 100 | 100 | 200 | < 100 | < 10 | 554 | [8 empty] | 1
| 1 | 2 | < 100 | 200 | < 100 | 100 | < 100 | < 10 | 496 | [8 empty] | 11
| 1 | 3 | < 100 | 100 | < 100 | 100 | < 100 | < 10 | 543 | [8 empty] | 16
| 1 | 4 | < 100 | 100 | 100 | 200 | 100 | < 10 | 550 | [8 empty] | 105
| 1 | 5 | < 100 | 200 | < 100 | 200 | 100 | < 10 | 526 | [8 empty] | 17
| 1 | 6 | < 100 | 100 | 300 | 100 | < 100 | < 10 | 591
| Mean |  |  | 133.3333 | 166.6667 | 150 | 100 |  | 543.3333
| Median |  |  | 100 | 100 | 150 | 100 |  | 546.5
| Std Dev. |  |  | 51.6398 | 115.4701 | 54.7723 | 0 |  | 31.532
| Rel.Std.Dev. |  |  | 38.73% | 69.28% | 36.51% | 0% |  | 5.803%
| PDM3 |  |  | 26.98% | 58.73% | 42.86% | -4.762% |  | 417.5%
| Z-Score (Absolute) |  | 0.94 | 0.4 | 0 | 0.67 | 0.67 | 1.67 | 7.04
| Analytical results for Au in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Au | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 2
2
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 12
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 18
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 313 (Certified Value 0.269 wt.%) | [16 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.18 | 2.19 | 2.20 | 2.24 | 2.25 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.256 | 0.29 | 0.26 | 0.26918 | 0.26 | 0.28 | 0.28 | 0.28 | 0.27 | 0.268 | 0.2568 | 0.27 | 0.30246 |  |  | 1
| 1 | 2 | 0.267 | 0.29 | 0.26 | 0.27461 | 0.27 | 0.27 | 0.28 | 0.29 | 0.27 | 0.246 | 0.2568 | 0.27 | 0.30067 |  |  | #N/A
| 1 | 3 | 0.265 | 0.29 | 0.26 | 0.26888 | 0.27 | 0.27 | 0.28 | 0.29 | 0.28 | 0.279 | 0.2568 | 0.27 | 0.30201 |  |  | 16
| 1 | 4 | 0.271 | 0.301 | 0.26 | 0.26514 | 0.27 | 0.27 | 0.29 | 0.29 | 0.29 | 0.279 | 0.2568 | 0.27 | 0.30067 |  |  | 0.26931
| 1 | 5 | 0.259 | 0.301 | 0.26 | 0.27204 | 0.27 | 0.26 | 0.28 | 0.28 | 0.25 | 0.268 | 0.2568 | 0.27 | 0.30123 |  |  | 118
| 1 | 6 | 0.263 | 0.29 | 0.26 | 0.26625 | 0.27 | 0.28 | 0.27 | 0.29 | 0.25 | 0.257 | 0.2568 | 0.28 | 0.3
| Mean |  | 0.2635 | 0.29367 | 0.26 | 0.26935 | 0.26833 | 0.27167 | 0.28 | 0.28667 | 0.26833 | 0.26617 | 0.2568 | 0.27167 | 0.30117
| Median |  | 0.264 | 0.29 | 0.26 | 0.26903 | 0.27 | 0.27 | 0.28 | 0.29 | 0.27 | 0.268 | 0.2568 | 0.27 | 0.30095
| Std Dev. |  | 0.0054314 | 0.0056804 | 0 | 0.0035341 | 0.0040825 | 0.0075277 | 0.0063246 | 0.005164 | 0.016021 | 0.012859 | 0 | 0.0040825 | 0.0009211
| Rel.Std.Dev. |  | 2.061% | 1.934% | 0% | 1.312% | 1.521% | 2.771% | 2.259% | 1.801% | 5.97% | 4.831% | 0% | 1.503% | 0.3058%
| PDM3 |  | -2.157% | 9.044% | -3.457% | 0.01462% | -0.3627% | 0.8751% | 3.969% | 6.445% | -0.3627% | -1.167% | -4.645% | 0.8751% | 11.83%
| Z-Score (Absolute) |  | 0.67 | 2.84 | 1.08 | 0 | 0.12 | 0.27 | 1.23 | 2 | 0.12 | 0.37 | 1.45 | 0.27 | 3.67
| Analytical results for Bi in OREAS 313 (Certified Value < 100 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | 100 | < 100 | < 10 | 41 | [8 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | 100 | < 100 | < 10 | 42 | [8 empty] | 34
| 1 | 3 | < 100 | < 100 | < 100 | 100 | < 100 | < 10 | 38 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | 100 | < 100 | < 10 | 37 | [8 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | 100 | < 100 | < 10 | 40 | [8 empty] | 119
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 39
| Mean |  |  |  |  | 100 |  |  | 39.5
| Median |  |  |  |  | 100 |  |  | 39.5
| Std Dev. |  |  |  |  | 0 |  |  | 1.8708
| Rel.Std.Dev. |  |  |  |  | 0% |  |  | 4.736%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 313 (Certified Value 1.91 wt.%) | [16 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 1.85 | 1.894 | 1.96 | 2.27 | 1.9098 | 1.89 | 1.89 | 1.91 | 1.9 | 1.97 | 1.96 | 1.93 | 2.1408 | 1.94 |  | 1
| 1 | 2 | 1.85 | 1.896 | 1.97 | 2.44 | 1.9286 | 1.92 | 1.85 | 1.9 | 1.9 | 1.95 | 1.95 | 1.91 | 2.1268 | 1.95 |  | #N/A
| 1 | 3 | 1.85 | 1.895 | 1.94 | 2.51 | 1.9237 | 1.93 | 1.84 | 1.91 | 1.92 | 1.92 | 1.96 | 1.92 | 2.1268 | 1.95 |  | 16
| 1 | 4 | 1.85 | 1.904 | 1.94 | 2.52 | 1.8794 | 1.9 | 1.86 | 1.9 | 1.91 | 1.99 | 1.93 | 1.91 | 2.1268 | 1.95 |  | 1.913
| 1 | 5 | 1.86 | 1.883 | 1.96 | 2.29 | 1.9169 | 1.93 | 1.82 | 1.91 | 1.92 | 1.94 | 1.94 | 1.92 | 2.1268 | 1.93 |  | 120
| 1 | 6 | 1.84 | 1.895 | 1.94 | 2.42 | 1.8929 | 1.92 | 1.91 | 1.89 | 1.88 | 1.97 | 1.95 | 1.91 | 2.0848 | 1.95
| Mean |  | 1.85 | 1.8945 | 1.9517 | 2.4083 | 1.9085 | 1.915 | 1.8617 | 1.9033 | 1.905 | 1.9567 | 1.9483 | 1.9167 | 2.1221 | 1.945
| Median |  | 1.85 | 1.895 | 1.95 | 2.43 | 1.9134 | 1.92 | 1.855 | 1.905 | 1.905 | 1.96 | 1.95 | 1.915 | 2.1268 | 1.95
| Std Dev. |  | 0.0063246 | 0.0067157 | 0.013292 | 0.10685 | 0.018947 | 0.016432 | 0.033116 | 0.008165 | 0.015166 | 0.025033 | 0.01169 | 0.008165 | 0.019128 | 0.0083666
| Rel.Std.Dev. |  | 0.3419% | 0.3545% | 0.681% | 4.437% | 0.9927% | 0.8581% | 1.779% | 0.429% | 0.7961% | 1.279% | 0.6% | 0.426% | 0.9013% | 0.4302%
| PDM3 |  | -3.295% | -0.9687% | 2.02% | 25.89% | -0.2344% | 0.1029% | -2.685% | -0.507% | -0.4198% | 2.281% | 1.845% | 0.19% | 10.93% | 1.671%
| Z-Score (Absolute) |  | 1.44 | 0.47 | 0.78 | 10.77 | 0.16 | 0.02 | 1.18 | 0.27 | 0.24 | 0.89 | 0.71 | 0.02 | 4.51 | 0.64
| Analytical results for Cd in OREAS 313 (Indicative Value 126 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.20 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | 129 | [13 empty] | 1
| 1 | 2 | < 100 | 129 | [13 empty] | 13
| 1 | 3 | < 100 | 121 | [13 empty] | 16
| 1 | 4 | < 100 | 117 | [13 empty] | 126.3333
| 1 | 5 | < 100 | 129 | [13 empty] | 19
| 1 | 6 | < 100 | 133
| Mean |  |  | 126.3333
| Median |  |  | 129
| Std Dev. |  |  | 6.0222
| Rel.Std.Dev. |  |  | 4.767%
| PDM3 |  |  | 2.665e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ce in OREAS 313 (Indicative Value 78 ppm) | [16 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 80 | < 80 | 80 | 80 | 50 | [10 empty] | 1
| 1 | 2 | < 80 | < 80 | 160 | 80 | 42 | [10 empty] | 14
| 1 | 3 | < 80 | < 80 | 80 | 160 | 45 | [10 empty] | 16
| 1 | 4 | < 80 | < 80 | 80 | 80 | 43 | [10 empty] | 78.483
| 1 | 5 | < 80 | < 80 | 80 | 80 | 42 | [10 empty] | 20
| 1 | 6 | < 80 | < 80 | 80 | 80 | 51
| Mean |  |  |  | 93.3333 | 93.3333 | 45.5
| Median |  |  |  | 80 | 80 | 44
| Std Dev. |  |  |  | 32.6599 | 32.6599 | 4.0373
| Rel.Std.Dev. |  |  |  | 34.99% | 34.99% | 8.873%
| PDM3 |  |  |  | 18.92% | 18.92% | -42.03%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 6.96 | 6.96 | 0
| Analytical results for Cl in OREAS 313 (Indicative Value 0.308 wt.%) | [16 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.20 | 2.25 | [13 empty] | wt.%
|  |  | BF*XRF | BF*XRF | [13 empty] | 3
3
| 1 | 1 | 0.662 | 0.0057 | [13 empty] | 1
| 1 | 2 | 0.569 | 0.0058 | [13 empty] | 15
| 1 | 3 | 0.574 | 0.0058 | [13 empty] | 16
| 1 | 4 | 0.642 | 0.0057 | [13 empty] | 0.30824
| 1 | 5 | 0.622 | 0.0061 | [13 empty] | 21
| 1 | 6 | 0.595 | 0.0071
| Mean |  | 0.61067 | 0.0060333
| Median |  | 0.6085 | 0.0058
| Std Dev. |  | 0.037607 | 0.00054283
| Rel.Std.Dev. |  | 6.158% | 8.997%
| PDM3 |  | 98.11% | -98.04%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 313 (Indicative Value 74 ppm) | [16 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.20 | 2.25 | [7 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 100 | < 100 | < 100 | 66 | 100 | 46 | [7 empty] | 1
| 1 | 2 | < 100 | < 100 | 100 | < 100 | < 100 | 44 | 100 | 47 | [7 empty] | 16
| 1 | 3 | < 100 | < 100 | 100 | < 100 | < 100 | 51 | 100 | 43 | [7 empty] | 16
| 1 | 4 | < 100 | < 100 | 100 | < 100 | < 100 | 51 | 100 | 45 | [7 empty] | 73.8347
| 1 | 5 | < 100 | < 100 | 100 | < 100 | < 100 | 44 | 100 | 45 | [7 empty] | 22
| 1 | 6 | < 100 | < 100 | 100 | < 100 | < 100 | 44 | 100 | 45
| Mean |  |  |  | 100 |  |  | 50 | 100 | 45.1667
| Median |  |  |  | 100 |  |  | 47.5 | 100 | 45
| Std Dev. |  |  |  | 0 |  |  | 8.5557 | 0 | 1.3292
| Rel.Std.Dev. |  |  |  | 0% |  |  | 17.11% | 0% | 2.943%
| PDM3 |  |  |  | 35.44% |  |  | -32.28% | 35.44% | -38.83%
| Z-Score (Absolute) |  | 0 | 0 | 391.77 | 0 | 0 | 1.35 | 391.77 | 37.87
| Analytical results for Cr2O3 in OREAS 313 (Certified Value 150 ppm) | [16 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 | 2.25 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | < 100 | 168.762 | 150 | 150 | 150 | < 150 | 200 | < 50 | < 100 | 292 | < 100 | 122.8 |  |  | 1
| 1 | 2 | < 100 | < 100 | 131.1948 | 150 | 150 | 150 | < 150 | 300 | < 50 | < 100 | 292 | 200 | 125.7 |  |  | #N/A
| 1 | 3 | < 100 | < 100 | 121.7507 | 150 | 150 | 150 | < 150 | 300 | < 50 | < 100 | 292 | < 100 | 130.1 |  |  | 16
| 1 | 4 | < 100 | < 100 | 149.8277 | 150 | 150 | 150 | < 150 | 200 | < 50 | < 100 | 146 | < 100 | 122.8 |  |  | 150.3948
| 1 | 5 | < 100 | < 100 | 150.0531 | 150 | 150 | 150 | < 150 | 100 | < 50 | < 100 | 292 | 200 | 122.8 |  |  | 121
| 1 | 6 | < 100 | < 100 | 121.8211 | 150 | 150 | 150 | < 150 | 200 | < 50 | < 100 | 146 | 200 | 122.8
| Mean |  |  |  | 140.5682 | 150 | 150 | 150 |  | 216.6667 |  |  | 243.3333 | 200 | 124.5
| Median |  |  |  | 140.5112 | 150 | 150 | 150 |  | 200 |  |  | 292 | 200 | 122.8
| Std Dev. |  |  |  | 18.783 | 0 | 0 | 0 |  | 75.2773 |  |  | 75.3941 | 0 | 2.9786
| Rel.Std.Dev. |  |  |  | 13.36% | 0% | 0% | 0% |  | 34.74% |  |  | 30.98% | 0% | 2.392%
| PDM3 |  |  |  | -6.534% | -0.2625% | -0.2625% | -0.2625% |  | 44.07% |  |  | 61.8% | 32.98% | -17.22%
| Z-Score (Absolute) |  | 0.97 | 0.97 | 0.2 | 0.27 | 0.27 | 0.27 | 0.67 | 1.19 | 1.3 | 0.97 | 1.54 | 0 | 0.01
| Analytical results for Cs in OREAS 313 (Indicative Value 22.2 ppm) | [16 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | < 100 | 27 | [13 empty] | 1
| 1 | 2 | < 100 | 20 | [13 empty] | 17
| 1 | 3 | < 100 | 22 | [13 empty] | 16
| 1 | 4 | < 100 | 19 | [13 empty] | 22.1667
| 1 | 5 | < 100 | 23 | [13 empty] | 23
| 1 | 6 | < 100 | 22
| Mean |  |  | 22.1667
| Median |  |  | 22
| Std Dev. |  |  | 2.7869
| Rel.Std.Dev. |  |  | 12.57%
| PDM3 |  |  | -1.443e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 313 (Certified Value 0.035 wt.%) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.25 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | OxBF*XRF | BF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.04 | 0.032199 | 0.038 | 0.039 | 0.047 | 0.037 | 0.0344 | 0.03 | 0.03 | 0.0337 | [5 empty] | 1
| 1 | 2 | 0.04 | 0.032883 | 0.037 | 0.034 | 0.032 | 0.038 | 0.0344 | 0.03 | 0.03 | 0.0344 | [5 empty] | 39
| 1 | 3 | 0.04 | 0.037058 | 0.041 | 0.036 | 0.033 | 0.039 | 0.0336 | 0.03 | 0.03 | 0.0336 | [5 empty] | 16
| 1 | 4 | 0.04 | 0.030062 | 0.037 | 0.038 | 0.031 | 0.037 | 0.0344 | 0.03 | 0.03 | 0.0335 | [5 empty] | 0.034665
| 1 | 5 | 0.05 | 0.03851 | 0.037 | 0.038 | 0.036 | 0.037 | 0.0344 | 0.03 | 0.03 | 0.0335 | [5 empty] | 122
| 1 | 6 | 0.04 | 0.025886 | 0.037 | 0.041 | 0.036 | 0.037 | 0.0344 | 0.03 | 0.03 | 0.0335
| Mean |  | 0.041667 | 0.032766 | 0.037833 | 0.037667 | 0.035833 | 0.0375 | 0.034267 | 0.03 | 0.03 | 0.0337
| Median |  | 0.04 | 0.032541 | 0.037 | 0.038 | 0.0345 | 0.037 | 0.0344 | 0.03 | 0.03 | 0.03355
| Std Dev. |  | 0.0040825 | 0.0046129 | 0.0016021 | 0.0024221 | 0.0058452 | 0.00083666 | 0.0003266 | 0 | 0 | 0.00035214
| Rel.Std.Dev. |  | 9.798% | 14.08% | 4.235% | 6.43% | 16.31% | 2.231% | 0.9531% | 0% | 0% | 1.045%
| PDM3 |  | 20.2% | -5.477% | 9.14% | 8.659% | 3.37% | 8.178% | -1.149% | -13.46% | -13.46% | -2.784%
| Z-Score (Absolute) |  | 1.75 | 0.6 | 0.74 | 0.7 | 0.21 | 0.65 | 0.21 | 1.32 | 1.32 | 0.35
| Analytical results for Dy in OREAS 313 (Indicative Value 75 ppm) | [16 empty] | Indicative
Dy | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 75 | [14 empty] | 1
| 1 | 2 | 75 | [14 empty] | 5
| 1 | 3 | 74 | [14 empty] | 16
| 1 | 4 | 75 | [14 empty] | 74.8333
| 1 | 5 | 76 | [14 empty] | 24
| 1 | 6 | 74
| Mean |  | 74.8333
| Median |  | 75
| Std Dev. |  | 0.75277
| Rel.Std.Dev. |  | 1.006%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Er | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 6
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 25
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Eu | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 7
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 26
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 313 (Certified Value 8.99 wt.%) | [16 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.18 | 2.19 | 2.20 | 2.24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 9.75 | 9.079 | 9.036 | 9.11 | 8.8596 | 8.764 | 9.05 | 9.036 | 9.036 | 9.09 | 8.864 | 8 | 9 | 9.093 | 8.89 | 1
| 1 | 2 | 9.74 | 9.111 | 9.107 | 9.05 | 8.9127 | 8.964 | 8.864 | 9.021 | 8.964 | 9.07 | 8.936 | 8.02 | 8.9 | 9.093 | 8.91 | #N/A
| 1 | 3 | 9.68 | 9.144 | 8.936 | 9 | 8.8342 | 9.007 | 8.864 | 8.979 | 9.093 | 8.91 | 8.864 | 8.05 | 8.9 | 9.093 | 8.93 | 16
| 1 | 4 | 9.78 | 9.122 | 9.179 | 9.06 | 8.782 | 8.893 | 8.921 | 8.993 | 9.036 | 9.19 | 8.878 | 7.92 | 8.95 | 9.05 | 8.93 | 8.9875
| 1 | 5 | 9.67 | 9.075 | 9.093 | 9.41 | 8.8287 | 8.979 | 8.821 | 9.007 | 9.107 | 9.02 | 8.85 | 7.93 | 9.02 | 9.079 | 8.87 | 123
| 1 | 6 | 9.77 | 9.117 | 9.093 | 9.17 | 8.7187 | 8.921 | 9.136 | 8.921 | 9.007 | 9.05 | 8.864 | 7.96 | 8.95 | 8.964 | 8.94
| Mean |  | 9.7317 | 9.108 | 9.074 | 9.1333 | 8.8226 | 8.9213 | 8.9427 | 8.9928 | 9.0405 | 9.055 | 8.876 | 7.98 | 8.9533 | 9.062 | 8.9117
| Median |  | 9.745 | 9.114 | 9.093 | 9.085 | 8.8314 | 8.9425 | 8.8925 | 9 | 9.036 | 9.06 | 8.864 | 7.98 | 8.95 | 9.086 | 8.92
| Std Dev. |  | 0.046224 | 0.026503 | 0.081638 | 0.14733 | 0.066504 | 0.087225 | 0.12373 | 0.040519 | 0.053279 | 0.091597 | 0.030699 | 0.051769 | 0.049666 | 0.050833 | 0.027142
| Rel.Std.Dev. |  | 0.475% | 0.291% | 0.8997% | 1.613% | 0.7538% | 0.9777% | 1.384% | 0.4506% | 0.5893% | 1.012% | 0.3459% | 0.6487% | 0.5547% | 0.5609% | 0.3046%
| PDM3 |  | 8.28% | 1.34% | 0.9622% | 1.622% | -1.835% | -0.7365% | -0.4991% | 0.05907% | 0.5894% | 0.7508% | -1.241% | -11.21% | -0.3804% | 0.8287% | -0.844%
| Z-Score (Absolute) |  | 6.15 | 0.96 | 0.67 | 1.17 | 1.42 | 0.59 | 0.42 | 0 | 0.4 | 0.52 | 0.97 | 8.43 | 0.33 | 0.58 | 0.68
| Analytical results for Ga in OREAS 313 (Indicative Value 99 ppm) | [16 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | 99 | [13 empty] | 1
| 1 | 2 | < 100 | 99 | [13 empty] | 21
| 1 | 3 | < 100 | 98 | [13 empty] | 16
| 1 | 4 | < 100 | 100 | [13 empty] | 98.8333
| 1 | 5 | < 100 | 99 | [13 empty] | 27
| 1 | 6 | < 100 | 98
| Mean |  |  | 98.8333
| Median |  |  | 99
| Std Dev. |  |  | 0.75277
| Rel.Std.Dev. |  |  | 0.7617%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Gd in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Gd | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 8
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 28
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [13 empty] | 1
| 1 | 2 | < 100 | < 10 | [13 empty] | 23
| 1 | 3 | < 100 | < 10 | [13 empty] | 16
| 1 | 4 | < 100 | < 10 | [13 empty] | < 10
| 1 | 5 | < 100 | < 10 | [13 empty] | 29
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for HfO2 in OREAS 313 (Certified Value < 100 ppm) | [16 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 200 | < 100 | < 100 | < 100 | < 100 | 15.3 | [9 empty] | 1
| 1 | 2 | < 200 | < 100 | < 100 | < 100 | < 100 | 17.7 | [9 empty] | #N/A
| 1 | 3 | < 200 | < 100 | < 100 | < 100 | < 100 | 15.3 | [9 empty] | 16
| 1 | 4 | < 200 | < 100 | < 100 | < 100 | 100 | 17.7 | [9 empty] | < 100
| 1 | 5 | < 200 | < 100 | < 100 | < 100 | 100 | 17.7 | [9 empty] | 124
| 1 | 6 | < 200 | < 100 | 100 | < 100 | < 100 | 15.3
| Mean |  |  |  | 100 |  | 100 | 16.5
| Median |  |  |  | 100 |  | 100 | 16.5
| Std Dev. | [5 empty] | 0 | 1.3145
| Rel.Std.Dev. | [5 empty] | 0% | 7.967%
PDM3
| Z-Score (Absolute) |  | 3.71 | 0.34 | 0.34 | 0.34 | 1.01 | 3.05
| Analytical results for Hg in OREAS 313 (Indicative Value < 100 ppm) | [16 empty] | Indicative
Hg | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.04 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | [14 empty] | 1
| 1 | 2 | < 100 | [14 empty] | 24
| 1 | 3 | < 100 | [14 empty] | 16
| 1 | 4 | < 100 | [14 empty] | < 100
| 1 | 5 | < 100 | [14 empty] | 30
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Ho | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 4
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 16
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 313 (Indicative Value < 100 ppm) | [16 empty] | Indicative
In | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.04 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | [14 empty] | 1
| 1 | 2 | < 100 | [14 empty] | 11
| 1 | 3 | < 100 | [14 empty] | 16
| 1 | 4 | < 100 | [14 empty] | < 100
| 1 | 5 | < 100 | [14 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 313 (Certified Value 2.43 wt.%) | [16 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 2.41 | 2.427 | 2.52 | 2.67 | 2.4283 | 2.42 | 2.48 | 2.44 | 2.44 | 2.41 | 2.177 | 2.44 | 2.4574 | 2.48 |  | 1
| 1 | 2 | 2.4 | 2.443 | 2.53 | 2.96 | 2.4318 | 2.37 | 2.43 | 2.41 | 2.45 | 2.4 | 2.158 | 2.42 | 2.4574 | 2.41 |  | #N/A
| 1 | 3 | 2.41 | 2.458 | 2.5 | 3.05 | 2.4107 | 2.42 | 2.44 | 2.41 | 2.47 | 2.35 | 2.152 | 2.42 | 2.4333 | 2.42 |  | 16
| 1 | 4 | 2.42 | 2.439 | 2.56 | 3 | 2.4094 | 2.44 | 2.45 | 2.41 | 2.47 | 2.41 | 2.14 | 2.43 | 2.4453 | 2.36 |  | 2.4291
| 1 | 5 | 2.43 | 2.431 | 2.56 | 3.03 | 2.414 | 2.48 | 2.41 | 2.42 | 2.47 | 2.39 | 2.158 | 2.42 | 2.4574 | 2.47 |  | 125
| 1 | 6 | 2.42 | 2.437 | 2.56 | 2.98 | 2.3989 | 2.4 | 2.5 | 2.4 | 2.47 | 2.39 | 2.138 | 2.44 | 2.4333 | 2.46
| Mean |  | 2.415 | 2.4392 | 2.5383 | 2.9483 | 2.4155 | 2.4217 | 2.4517 | 2.415 | 2.4617 | 2.3917 | 2.1538 | 2.4283 | 2.4474 | 2.4333
| Median |  | 2.415 | 2.438 | 2.545 | 2.99 | 2.4123 | 2.42 | 2.445 | 2.41 | 2.47 | 2.395 | 2.155 | 2.425 | 2.4513 | 2.44
| Std Dev. |  | 0.010488 | 0.010852 | 0.025626 | 0.1402 | 0.012378 | 0.037103 | 0.033116 | 0.013784 | 0.013292 | 0.022286 | 0.014261 | 0.0098319 | 0.011849 | 0.045461
| Rel.Std.Dev. |  | 0.4343% | 0.4449% | 1.01% | 4.755% | 0.5124% | 1.532% | 1.351% | 0.5708% | 0.5399% | 0.9318% | 0.6621% | 0.4049% | 0.4842% | 1.868%
| PDM3 |  | -0.5814% | 0.4134% | 4.496% | 21.37% | -0.5603% | -0.307% | 0.928% | -0.5814% | 1.34% | -1.542% | -11.33% | -0.03253% | 0.7503% | 0.1733%
| Z-Score (Absolute) |  | 0.66 | 0.35 | 4.48 | 21.58 | 0.64 | 0.38 | 0.87 | 0.66 | 1.29 | 1.63 | 11.55 | 0.1 | 0.69 | 0.1
| Analytical results for La in OREAS 313 (Indicative Value 62 ppm) | [16 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 90 | < 90 | 90 | < 90 | 24 | [10 empty] | 1
| 1 | 2 | < 90 | < 90 | 170 | < 90 | 24 | [10 empty] | 12
| 1 | 3 | < 90 | < 90 | < 90 | < 90 | 24 | [10 empty] | 16
| 1 | 4 | < 90 | < 90 | 170 | < 90 | 24 | [10 empty] | 61.7385
| 1 | 5 | < 90 | < 90 | 90 | < 90 | 24 | [10 empty] | 18
| 1 | 6 | < 90 | < 90 | < 90 | < 90 | 24
| Mean |  |  |  | 130 |  | 24
| Median |  |  |  | 130 |  | 24
| Std Dev. |  |  |  | 46.188 |  | 0
| Rel.Std.Dev. |  |  |  | 35.53% |  | 0%
| PDM3 |  |  |  | 110.6% |  | -61.13%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Lu in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Lu | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 5
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 19
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 313 (Certified Value 1.76 wt.%) | [16 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 2.24 | 1.69 | 1.6 | 1.68 | 1.8565 | 1.73 | 1.82 | 1.82 | 1.86 | 1.71 | 1.71 | 1.79 | 1.691 | 1.81 |  | 1
| 1 | 2 | 2.21 | 1.698 | 1.61 | 1.9 | 1.8565 | 1.78 | 1.78 | 1.84 | 1.82 | 1.73 | 1.69 | 1.77 | 1.675 | 1.78 |  | #N/A
| 1 | 3 | 2.22 | 1.693 | 1.58 | 2 | 1.8475 | 1.8 | 1.79 | 1.82 | 1.85 | 1.68 | 1.71 | 1.77 | 1.675 | 1.8 |  | 16
| 1 | 4 | 2.24 | 1.652 | 1.61 | 2 | 1.8264 | 1.78 | 1.8 | 1.84 | 1.84 | 1.72 | 1.69 | 1.77 | 1.675 | 1.8 |  | 1.7631
| 1 | 5 | 2.25 | 1.674 | 1.59 | 1.92 | 1.8351 | 1.79 | 1.76 | 1.82 | 1.87 | 1.71 | 1.68 | 1.77 | 1.691 | 1.79 |  | 126
| 1 | 6 | 2.24 | 1.673 | 1.61 | 1.39 | 1.8188 | 1.76 | 1.82 | 1.8 | 1.85 | 1.74 | 1.69 | 1.77 | 1.625 | 1.79
| Mean |  | 2.2333 | 1.68 | 1.6 | 1.815 | 1.8401 | 1.7733 | 1.795 | 1.8233 | 1.8483 | 1.715 | 1.695 | 1.7733 | 1.672 | 1.795
| Median |  | 2.24 | 1.682 | 1.605 | 1.91 | 1.8413 | 1.78 | 1.795 | 1.82 | 1.85 | 1.715 | 1.69 | 1.77 | 1.675 | 1.795
| Std Dev. |  | 0.015055 | 0.0171 | 0.012649 | 0.23898 | 0.015878 | 0.025033 | 0.023452 | 0.015055 | 0.017224 | 0.020736 | 0.012247 | 0.008165 | 0.024323 | 0.010488
| Rel.Std.Dev. |  | 0.6741% | 1.018% | 0.7906% | 13.17% | 0.8629% | 1.412% | 1.307% | 0.8257% | 0.9319% | 1.209% | 0.7226% | 0.4604% | 1.455% | 0.5843%
| PDM3 |  | 26.67% | -4.712% | -9.249% | 2.946% | 4.371% | 0.5823% | 1.811% | 3.418% | 4.836% | -2.726% | -3.861% | 0.5823% | -5.165% | 1.811%
| Z-Score (Absolute) |  | 5.04 | 1.17 | 2.07 | 0.35 | 0.63 | 0.12 | 0.12 | 0.44 | 0.72 | 0.78 | 1 | 0.12 | 1.26 | 0.12
| Analytical results for MnO in OREAS 313 (Certified Value 0.246 wt.%) | [16 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.24 | 0.174 | 0.25 | 0.27 | 0.25629 | 0.245 | 0.258 | 0.245 | 0.245 | 0.239 | 0.351 | 0.24 | 0.297 | 0.25 |  | 1
| 1 | 2 | 0.24 | 0.188 | 0.24 | 0.34 | 0.25617 | 0.258 | 0.245 | 0.245 | 0.245 | 0.241 | 0.35 | 0.24 | 0.297 | 0.25 |  | #N/A
| 1 | 3 | 0.24 | 0.18 | 0.25 | 0.29 | 0.2541 | 0.258 | 0.245 | 0.245 | 0.245 | 0.243 | 0.352 | 0.24 | 0.3099 | 0.25 |  | 16
| 1 | 4 | 0.24 | 0.189 | 0.25 | 0.27 | 0.2512 | 0.245 | 0.245 | 0.245 | 0.245 | 0.241 | 0.339 | 0.25 | 0.297 | 0.25 |  | 0.24639
| 1 | 5 | 0.24 | 0.205 | 0.25 | 0.33 | 0.25158 | 0.258 | 0.245 | 0.245 | 0.245 | 0.238 | 0.351 | 0.24 | 0.297 | 0.25 |  | 127
| 1 | 6 | 0.24 | 0.203 | 0.24 | 0.28 | 0.25043 | 0.258 | 0.258 | 0.232 | 0.245 | 0.239 | 0.35 | 0.24 | 0.297 | 0.25
| Mean |  | 0.24 | 0.18983 | 0.24667 | 0.29667 | 0.2533 | 0.25367 | 0.24933 | 0.24283 | 0.245 | 0.24017 | 0.34883 | 0.24167 | 0.29915 | 0.25
| Median |  | 0.24 | 0.1885 | 0.25 | 0.285 | 0.25284 | 0.258 | 0.245 | 0.245 | 0.245 | 0.24 | 0.3505 | 0.24 | 0.297 | 0.25
| Std Dev. |  | 0 | 0.012287 | 0.005164 | 0.030768 | 0.0025853 | 0.0067132 | 0.0067132 | 0.0053072 | 3.0405e-17 | 0.0018348 | 0.0048751 | 0.0040825 | 0.0052664 | 0
| Rel.Std.Dev. |  | 0% | 6.472% | 2.094% | 10.37% | 1.021% | 2.646% | 2.692% | 2.186% | 1.241e-14% | 0.764% | 1.398% | 1.689% | 1.76% | 0%
| PDM3 |  | -2.593% | -22.95% | 0.1127% | 20.41% | 2.803% | 2.954% | 1.195% | -1.443% | -0.5637% | -2.525% | 41.58% | -1.917% | 21.41% | 1.466%
| Z-Score (Absolute) |  | 0.9 | 6.41 | 0.16 | 5.33 | 0.57 | 0.64 | 0.16 | 0.55 | 0.31 | 0.88 | 11.06 | 0.71 | 5.6 | 0.2
| Analytical results for Mo in OREAS 313 (Indicative Value 40.1 ppm) | [16 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 1
1
| 1 | 1 | < 100 | < 50 | < 50 | 80 | < 50 | 14 | 34 | [8 empty] | 1
| 1 | 2 | < 100 | 50 | < 50 | 50 | < 50 | 15 | 34 | [8 empty] | 14
| 1 | 3 | < 100 | < 50 | < 50 | 70 | < 50 | 16 | 35 | [8 empty] | 16
| 1 | 4 | < 100 | 50 | < 50 | 80 | < 50 | 16 | 34 | [8 empty] | 40.0833
| 1 | 5 | < 100 | 60 | < 50 | 90 | < 50 | 15 | 35 | [8 empty] | 20
| 1 | 6 | < 100 | < 50 | < 50 | 60 | < 50 | 15 | 34
| Mean |  |  | 53.3333 |  | 71.6667 |  | 15.1667 | 34.3333
| Median |  |  | 50 |  | 75 |  | 15 | 34
| Std Dev. |  |  | 5.7735 |  | 14.7196 |  | 0.75277 | 0.5164
| Rel.Std.Dev. |  |  | 10.83% |  | 20.54% |  | 4.963% | 1.504%
| PDM3 |  |  | 33.06% |  | 78.79% |  | -62.16% | -14.35%
| Z-Score (Absolute) |  | 1.13 | 0.35 | 0.67 | 2.7 | 0.67 | 1.38 | 0
| Analytical results for Na2O in OREAS 313 (Indicative Value 0.222 wt.%) | [16 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.18 | 2.24 | [8 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 3
3
| 1 | 1 | 0.62 | 0.077 | 0.214 | < 0.01 | 0.99382 | 0.17 | 0.02 | [8 empty] | 1
| 1 | 2 | 0.59 | 0.077 | 0.214 | < 0.01 | 1.0027 | 0.28 | 0.01 | [8 empty] | 15
| 1 | 3 | 0.63 | 0.075 | 0.179 | < 0.01 | 0.99367 | 0.17 | 0.02 | [8 empty] | 16
| 1 | 4 | 0.64 | 0.077 | 0.216 | < 0.01 | 0.97763 | 0.18 | 0.01 | [8 empty] | 0.22187
| 1 | 5 | 0.63 | 0.076 | 0.21 | < 0.01 | 0.98848 | 0.19 | 0.02 | [8 empty] | 21
| 1 | 6 | 0.64 | 0.075 | 0.202 | < 0.01 | 0.97363 | 0.2 | 0.01
| Mean |  | 0.625 | 0.076167 | 0.20583 |  | 0.98832 | 0.19833 | 0.015
| Median |  | 0.63 | 0.0765 | 0.212 |  | 0.99107 | 0.185 | 0.015
| Std Dev. |  | 0.018708 | 0.00098319 | 0.014063 |  | 0.010917 | 0.041673 | 0.0054772
| Rel.Std.Dev. |  | 2.993% | 1.291% | 6.832% |  | 1.105% | 21.01% | 36.51%
| PDM3 |  | 181.7% | -65.67% | -7.229% |  | 345.4% | -10.61% | -93.24%
| Z-Score (Absolute) |  | 1.57 | 0.45 | 0.03 | 0.71 | 2.91 | 0 | 0.67
| Analytical results for Nb in OREAS 313 (Indicative Value 60 ppm) | [16 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 60 | < 50 | 60 | < 50 | 46 | [10 empty] | 1
| 1 | 2 | 70 | < 50 | 70 | < 50 | 46 | [10 empty] | 16
| 1 | 3 | 70 | < 50 | 80 | < 50 | 46 | [10 empty] | 16
| 1 | 4 | 70 | < 50 | 60 | < 50 | 46 | [10 empty] | 60.3333
| 1 | 5 | 70 | < 50 | 70 | < 50 | 46 | [10 empty] | 22
| 1 | 6 | 70 | < 50 | 60 | < 50 | 46
| Mean |  | 68.3333 |  | 66.6667 |  | 46
| Median |  | 70 |  | 65 |  | 46
| Std Dev. |  | 4.0825 |  | 8.165 |  | 0
| Rel.Std.Dev. |  | 5.974% |  | 12.25% |  | 0%
| PDM3 |  | 13.26% |  | 10.5% |  | -23.76%
| Z-Score (Absolute) |  | 0.72 | 0.67 | 0.66 | 0.67 | 0
| Analytical results for Nd in OREAS 313 (Indicative Value 103 ppm) | [16 empty] | Indicative
Nd | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 105 | [14 empty] | 1
| 1 | 2 | 104 | [14 empty] | 7
| 1 | 3 | 118 | [14 empty] | 16
| 1 | 4 | 97 | [14 empty] | 103
| 1 | 5 | 98 | [14 empty] | 23
| 1 | 6 | 96
| Mean |  | 103
| Median |  | 101
| Std Dev. |  | 8.2462
| Rel.Std.Dev. |  | 8.006%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 313 (Indicative Value 50 ppm) | [16 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.20 | 2.25 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 196 | < 50 | < 50 | 100 | < 50 | 55 | < 100 | < 100 | 12 | [5 empty] | 1
| 1 | 2 | < 100 | < 196 | < 50 | < 50 | 130 | < 50 | 39 | < 100 | < 100 | 13 | [5 empty] | 18
| 1 | 3 | < 100 | < 196 | 50 | < 50 | 110 | < 50 | 55 | < 100 | < 100 | 12 | [5 empty] | 16
| 1 | 4 | < 100 | < 196 | < 50 | < 50 | 60 | < 50 | 31 | < 100 | < 100 | 14 | [5 empty] | 50.4966
| 1 | 5 | < 100 | < 196 | < 50 | < 50 | 90 | < 50 | 71 | < 100 | < 100 | 12 | [5 empty] | 24
| 1 | 6 | < 100 | < 196 | 50 | < 50 | 110 | < 50 | 86 | < 100 | < 100 | 11
| Mean |  |  |  | 50 |  | 100 |  | 56.1667 |  |  | 12.3333
| Median |  |  |  | 50 |  | 105 |  | 55 |  |  | 12
| Std Dev. |  |  |  | 0 |  | 23.6643 |  | 20.2031 |  |  | 1.0328
| Rel.Std.Dev. |  |  |  | 0% |  | 23.66% |  | 35.97% |  |  | 8.374%
| PDM3 |  |  |  | -0.9834% |  | 98.03% |  | 11.23% |  |  | -75.58%
| Z-Score (Absolute) |  | 0 | 1.56 | 0.54 | 0.81 | 1.62 | 0.81 | 0.2 | 0 | 0 | 1.22
| Analytical results for P2O5 in OREAS 313 (Certified Value 0.071 wt.%) | [16 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.066 | 0.075 | 0.085 | 0.18 | 0.063989 | 0.07 | 0.06 | 0.07 | 0.08 | 0.07 | 0.02 | 0.07 | 0.0687 | 0.06 |  | 1
| 1 | 2 | 0.066 | 0.074 | 0.076 | 0.18 | 0.065925 | 0.07 | 0.06 | 0.08 | 0.08 | 0.07 | 0.013 | 0.08 | 0.0687 | 0.07 |  | #N/A
| 1 | 3 | 0.064 | 0.073 | 0.075 | 0.14 | 0.063919 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.014 | 0.07 | 0.0687 | 0.08 |  | 16
| 1 | 4 | 0.066 | 0.073 | 0.077 | 0.19 | 0.061944 | 0.07 | 0.06 | 0.07 | 0.08 | 0.07 | 0.012 | 0.08 | 0.0687 | 0.08 |  | 0.070906
| 1 | 5 | 0.064 | 0.077 | 0.073 | 0.12 | 0.063022 | 0.07 | 0.06 | 0.07 | 0.08 | 0.07 | 0.015 | 0.07 | 0.0687 | 0.08 |  | 128
| 1 | 6 | 0.064 | 0.071 | 0.08 | 0.18 | 0.06494 | 0.07 | 0.06 | 0.07 | 0.08 | 0.08 | 0.02 | 0.07 | 0.0687 | 0.07
| Mean |  | 0.065 | 0.073833 | 0.077667 | 0.165 | 0.063957 | 0.07 | 0.061667 | 0.071667 | 0.08 | 0.071667 | 0.015667 | 0.073333 | 0.0687 | 0.073333
| Median |  | 0.065 | 0.0735 | 0.0765 | 0.18 | 0.063954 | 0.07 | 0.06 | 0.07 | 0.08 | 0.07 | 0.0145 | 0.07 | 0.0687 | 0.075
| Std Dev. |  | 0.0010954 | 0.0020412 | 0.004274 | 0.028107 | 0.0013977 | 0 | 0.0040825 | 0.0040825 | 0 | 0.0040825 | 0.0035024 | 0.005164 | 0 | 0.008165
| Rel.Std.Dev. |  | 1.685% | 2.765% | 5.503% | 17.03% | 2.185% | 0% | 6.62% | 5.696% | 0% | 5.696% | 22.36% | 7.042% | 0% | 11.13%
| PDM3 |  | -8.329% | 4.129% | 9.535% | 132.7% | -9.801% | -1.278% | -13.03% | 1.073% | 12.83% | 1.073% | -77.91% | 3.423% | -3.111% | 3.423%
| Z-Score (Absolute) |  | 1.01 | 0.33 | 0.91 | 14.11 | 1.17 | 0.25 | 1.51 | 0 | 1.26 | 0 | 8.47 | 0.25 | 0.44 | 0.25
| Analytical results for Pb in OREAS 313 (Certified Value 3.86 wt.%) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  |  | 2
3
| 1 | 1 | 3.853 | 3.77 | 3.5422 | 3.75 | 4.02 | 3.9 | 3.94 | 3.935 | 3.9 | 3.73 | 3.83 | 3.8344 |  |  |  | 1
| 1 | 2 | 3.955 | 3.68 | 3.5601 | 3.81 | 3.93 | 3.85 | 3.93 | 3.898 | 3.93 | 3.7 | 3.87 | 3.8435 |  |  |  | 40
| 1 | 3 | 4.122 | 3.88 | 3.5126 | 3.84 | 3.93 | 3.88 | 3.99 | 3.79 | 3.9 | 3.74 | 3.85 | 3.8454 |  |  |  | 16
| 1 | 4 | 3.908 | 3.63 | 3.5082 | 3.81 | 3.95 | 3.88 | 3.98 | 3.934 | 3.87 | 3.73 | 3.8 | 3.8333 |  |  |  | 3.8597
| 1 | 5 | 3.936 | 3.89 | 3.5393 | 3.85 | 3.9 | 3.89 | 4.01 | 3.869 | 3.88 | 3.73 | 3.81 | 3.8421 |  |  |  | 129
| 1 | 6 | 3.945 | 3.71 | 3.4942 | 3.79 | 4.07 | 3.84 | 3.95 | 3.854 | 3.89 | 3.69 | 3.84 | 3.8024
| Mean |  | 3.9532 | 3.76 | 3.5261 | 3.8083 | 3.9667 | 3.8733 | 3.9667 | 3.88 | 3.895 | 3.72 | 3.8333 | 3.8335
| Median |  | 3.9405 | 3.74 | 3.5259 | 3.81 | 3.94 | 3.88 | 3.965 | 3.8835 | 3.895 | 3.73 | 3.835 | 3.8383
| Std Dev. |  | 0.090491 | 0.10696 | 0.024938 | 0.036009 | 0.064704 | 0.023381 | 0.031411 | 0.055067 | 0.020736 | 0.02 | 0.02582 | 0.016023
| Rel.Std.Dev. |  | 2.289% | 2.845% | 0.7073% | 0.9455% | 1.631% | 0.6036% | 0.7919% | 1.419% | 0.5324% | 0.5376% | 0.6736% | 0.418%
| PDM3 |  | 2.423% | -2.582% | -8.642% | -1.33% | 2.772% | 0.3542% | 2.772% | 0.527% | 0.9156% | -3.618% | -0.6821% | -0.6774%
| Z-Score (Absolute) |  | 0.97 | 0.91 | 3.19 | 0.44 | 1.1 | 0.19 | 1.1 | 0.26 | 0.4 | 1.3 | 0.2 | 0.19
| Analytical results for Pr in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Pr | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 19
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 25
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 313 (Indicative Value 83 ppm) | [16 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | 60 | 90 | 100 | 60 | 163 | [9 empty] | 1
| 1 | 2 | < 100 | 70 | 70 | 110 | 70 | 163 | [9 empty] | 20
| 1 | 3 | < 100 | 70 | 90 | 100 | 100 | 160 | [9 empty] | 16
| 1 | 4 | < 100 | 70 | 80 | 90 | 80 | 161 | [9 empty] | 83.3333
| 1 | 5 | < 100 | 70 | 50 | 120 | 80 | 162 | [9 empty] | 26
| 1 | 6 | < 100 | 70 | 130 | 80 | 90 | 161
| Mean |  |  | 68.3333 | 85 | 100 | 80 | 161.6667
| Median |  |  | 70 | 85 | 100 | 80 | 161.5
| Std Dev. |  |  | 4.0825 | 26.6458 | 14.1421 | 14.1421 | 1.2111
| Rel.Std.Dev. |  |  | 5.974% | 31.35% | 14.14% | 17.68% | 0.7491%
| PDM3 |  |  | -18% | 2% | 20% | -4% | 94%
| Z-Score (Absolute) |  | 1.38 | 0.6 | 0.11 | 0.75 | 0.11 | 3.37
| Analytical results for S in OREAS 313 (Certified Value 4.59 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.18 | 2.20 | 2.25 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 2
3
| 1 | 1 | > 2 | 4.317 | 4.63 | 4.74 | 4.52 | 4.6 | 4.7093 | 4.65 | 3.4672 | [6 empty] | 1
| 1 | 2 | > 2 | 4.385 | 4.57 | 4.66 | 4.54 | 4.67 | 4.7173 | 4.65 | 3.471 | [6 empty] | #N/A
| 1 | 3 | > 2 | 4.269 | 4.45 | 4.71 | 4.5 | 4.68 | 4.6961 | 4.68 | 3.4744 | [6 empty] | 16
| 1 | 4 | > 2 | 4.377 | 4.65 | 4.68 | 4.47 | 4.67 | 4.6729 | 4.75 | 3.4771 | [6 empty] | 4.5904
| 1 | 5 | > 2 | 4.333 | 4.68 | 4.64 | 4.47 | 4.68 | 4.6452 | 4.71 | 3.4671 | [6 empty] | 130
| 1 | 6 | > 2 | 4.321 | 4.38 | 4.79 | 4.58 | 4.66 | 4.6244 | 4.67 | 3.4561
| Mean |  |  | 4.3337 | 4.56 | 4.7033 | 4.5133 | 4.66 | 4.6775 | 4.685 | 3.4688
| Median |  |  | 4.327 | 4.6 | 4.695 | 4.51 | 4.67 | 4.6845 | 4.675 | 3.4691
| Std Dev. |  |  | 0.04274 | 0.12 | 0.055377 | 0.04274 | 0.030332 | 0.036947 | 0.038859 | 0.0073733
| Rel.Std.Dev. |  |  | 0.9862% | 2.632% | 1.177% | 0.947% | 0.6509% | 0.7899% | 0.8294% | 0.2126%
| PDM3 |  |  | -5.593% | -0.662% | 2.46% | -1.679% | 1.516% | 1.898% | 2.061% | -24.43%
| Z-Score (Absolute) |  | NA | 2.21 | 0.4 | 0.75 | 0.77 | 0.4 | 0.54 | 0.6 | 9.14
| Analytical results for Sb in OREAS 313 (Indicative Value < 50 ppm) | [16 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | 60 | < 50 | 60 | < 50 | < 10 | 50 | [8 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 51 | [8 empty] | 21
| 1 | 3 | < 100 | 70 | < 50 | < 50 | < 50 | < 10 | 50 | [8 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | 50 | < 50 | < 10 | 50 | [8 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 49 | [8 empty] | 27
| 1 | 6 | < 100 | 50 | < 50 | < 50 | < 50 | < 10 | 49
| Mean |  |  | 60 |  | 55 |  |  | 49.8333
| Median |  |  | 60 |  | 55 |  |  | 50
| Std Dev. |  |  | 10 |  | 7.0711 |  |  | 0.75277
| Rel.Std.Dev. |  |  | 16.67% |  | 12.86% |  |  | 1.511%
PDM3
| Z-Score (Absolute) |  | 1.01 | 0.51 | 0.67 | 0 | 0.67 | 2.02 | 1
| Analytical results for Sc in OREAS 313 (Indicative Value 16.8 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | < 100 | 18 | [13 empty] | 1
| 1 | 2 | < 100 | 16 | [13 empty] | 22
| 1 | 3 | < 100 | 16 | [13 empty] | 16
| 1 | 4 | < 100 | 18 | [13 empty] | 16.8333
| 1 | 5 | < 100 | 17 | [13 empty] | 28
| 1 | 6 | < 100 | 16
| Mean |  |  | 16.8333
| Median |  |  | 16.5
| Std Dev. |  |  | 0.98319
| Rel.Std.Dev. |  |  | 5.841%
| PDM3 |  |  | 1.998e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Se in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [13 empty] | 1
| 1 | 2 | < 100 | < 10 | [13 empty] | 2
| 1 | 3 | < 100 | < 10 | [13 empty] | 16
| 1 | 4 | < 100 | < 10 | [13 empty] | < 10
| 1 | 5 | < 100 | < 10 | [13 empty] | 29
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for SiO2 in OREAS 313 (Certified Value 57.61 wt.%) | [16 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 57.54 | 56.992 | 58.78 | 54.29 | 57.7549 | 57.1 | 57.8 | 57.4 | 58.1 | 58.03 | 50.63 | 57.59 | 58.657 | 57.26 |  | 1
| 1 | 2 | 57.52 | 56.813 | 57.24 | 52.89 | 57.7921 | 58.3 | 56.8 | 57.2 | 58 | 57.71 | 50.65 | 57.35 | 57.951 | 57.47 |  | #N/A
| 1 | 3 | 57.85 | 56.812 | 57.21 | 53.19 | 57.3396 | 58.4 | 57 | 57.1 | 58.4 | 56.7 | 50.05 | 57.41 | 58.443 | 57.51 |  | 16
| 1 | 4 | 57.93 | 57.157 | 57.63 | 52.26 | 57.1461 | 57.9 | 57.2 | 57.3 | 58.2 | 58.08 | 50.05 | 57.51 | 58.635 | 57.51 |  | 57.6097
| 1 | 5 | 57.99 | 56.965 | 58.48 | 53.15 | 57.4089 | 58.3 | 56.3 | 57.1 | 58.6 | 57.34 | 50.08 | 57.67 | 58.25 | 57.33 |  | 131
| 1 | 6 | 57.42 | 57.102 | 58.27 | 53.07 | 56.8796 | 57.9 | 58.1 | 56.6 | 58.1 | 57.53 | 50.12 | 57.51 | 57.951 | 57.28
| Mean |  | 57.7083 | 56.9735 | 57.935 | 53.1417 | 57.3869 | 57.9833 | 57.2 | 57.1167 | 58.2333 | 57.565 | 50.2633 | 57.5067 | 58.3145 | 57.3933
| Median |  | 57.695 | 56.9785 | 57.95 | 53.11 | 57.3743 | 58.1 | 57.1 | 57.15 | 58.15 | 57.62 | 50.1 | 57.51 | 58.3465 | 57.4
| Std Dev. |  | 0.2431 | 0.14314 | 0.66708 | 0.65831 | 0.3515 | 0.48339 | 0.6603 | 0.27869 | 0.22509 | 0.5104 | 0.29296 | 0.11622 | 0.31779 | 0.11639
| Rel.Std.Dev. |  | 0.4213% | 0.2512% | 1.151% | 1.239% | 0.6125% | 0.8337% | 1.154% | 0.4879% | 0.3865% | 0.8867% | 0.5829% | 0.2021% | 0.545% | 0.2028%
| PDM3 |  | 0.1712% | -1.104% | 0.5646% | -7.756% | -0.3868% | 0.6485% | -0.7112% | -0.8558% | 1.082% | -0.07762% | -12.75% | -0.1789% | 1.223% | -0.3756%
| Z-Score (Absolute) |  | 0.43 | 0.79 | 0.81 | 7.17 | 0.11 | 0.89 | 0.42 | 0.56 | 1.3 | 0.19 | 11.97 | 0.09 | 1.44 | 0.09
| Analytical results for Sm in OREAS 313 (Indicative Value 24 ppm) | [16 empty] | Indicative
Sm | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | 26 | [14 empty] | 1
| 1 | 2 | 26 | [14 empty] | 24
| 1 | 3 | 21 | [14 empty] | 16
| 1 | 4 | 26 | [14 empty] | 24
| 1 | 5 | 25 | [14 empty] | 30
| 1 | 6 | 20
| Mean |  | 24
| Median |  | 25.5
| Std Dev. |  | 2.7568
| Rel.Std.Dev. |  | 11.49%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 313 (Certified Value < 50 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | [7 empty] | 1
| 1 | 2 | < 100 | < 160 | < 50 | < 50 | 70 | < 50 | < 100 | < 10 | [7 empty] | 31
| 1 | 3 | < 100 | < 160 | < 50 | < 50 | 60 | < 50 | < 100 | < 10 | [7 empty] | 16
| 1 | 4 | < 100 | < 160 | < 50 | < 50 | 70 | < 50 | < 100 | < 10 | [7 empty] | < 50
| 1 | 5 | < 100 | < 160 | < 50 | < 50 | 60 | 50 | < 100 | < 10 | [7 empty] | 132
| 1 | 6 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10
| Mean | [5 empty] | 65 | 50
| Median | [5 empty] | 65 | 50
| Std Dev. | [5 empty] | 5.7735
| Rel.Std.Dev. | [5 empty] | 8.882%
PDM3
| Z-Score (Absolute) |  | 0.53 | 1.98 | 0.74 | 0.74 | 0.61 | 0.53 | 0.53 | 1.75
| Analytical results for SrO in OREAS 313 (Indicative Value 76 ppm) | [16 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.18 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 120 | 120 | 120 | < 120 | 50 | < 100 | 60.3 | [6 empty] | 1
| 1 | 2 | < 100 | < 50 | < 120 | < 120 | 120 | < 120 | 60 | < 100 | 61.5 | [6 empty] | 10
| 1 | 3 | < 100 | < 50 | < 120 | < 120 | 120 | < 120 | < 50 | < 100 | 60.3 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 120 | 120 | 120 | < 120 | 60 | < 100 | 59.1 | [6 empty] | 75.9613
| 1 | 5 | < 100 | < 50 | < 120 | < 120 | 120 | 120 | 50 | < 100 | 59.1 | [6 empty] | 16
| 1 | 6 | < 100 | < 50 | < 120 | 120 | < 120 | < 120 | 50 | < 100 | 57.9
| Mean |  |  |  |  | 120 | 120 | 120 | 54 |  | 59.7
| Median |  |  |  |  | 120 | 120 | 120 | 50 |  | 59.7
| Std Dev. |  |  |  |  | 0 | 0 |  | 5.4772 |  | 1.2586
| Rel.Std.Dev. |  |  |  |  | 0% | 0% |  | 10.14% |  | 2.108%
| PDM3 |  |  |  |  | 57.98% | 57.98% | 57.98% | -28.91% |  | -21.41%
| Z-Score (Absolute) |  | 0.62 | 2.34 | 0 | 2.02 | 3.37 | 0.67 | 0.68 | 0.62 | 0.04
| Analytical results for Ta in OREAS 313 (Indicative Value 12 ppm) | [16 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | < 100 | 12 | [13 empty] | 1
| 1 | 2 | < 100 | 12 | [13 empty] | 11
| 1 | 3 | < 100 | 12 | [13 empty] | 16
| 1 | 4 | < 100 | 12 | [13 empty] | 12
| 1 | 5 | < 100 | 12 | [13 empty] | 17
| 1 | 6 | < 100 | 12
| Mean |  |  | 12
| Median |  |  | 12
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Tb in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Tb | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 12
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 18
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 313 (Indicative Value < 100 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.04 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | [14 empty] | 1
| 1 | 2 | < 100 | [14 empty] | 13
| 1 | 3 | < 100 | [14 empty] | 16
| 1 | 4 | < 100 | [14 empty] | < 100
| 1 | 5 | < 100 | [14 empty] | 19
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 313 (Indicative Value 516 ppm) | [16 empty] | Indicative
Th | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 513 | [14 empty] | 1
| 1 | 2 | 516 | [14 empty] | 14
| 1 | 3 | 520 | [14 empty] | 16
| 1 | 4 | 517 | [14 empty] | 515.5
| 1 | 5 | 514 | [14 empty] | 20
| 1 | 6 | 513
| Mean |  | 515.5
| Median |  | 515
| Std Dev. |  | 2.7386
| Rel.Std.Dev. |  | 0.5313%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 313 (Certified Value 0.407 wt.%) | [16 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.37 | 0.428 | 0.42 | < 0.01 | 0.40975 | 0.41 | 0.41 | 0.42 | 0.4 | 0.39 | 0.37 | 0.41 | 0.417 | 0.4 |  | 1
| 1 | 2 | 0.36 | 0.429 | 0.41 | < 0.01 | 0.4036 | 0.41 | 0.4 | 0.42 | 0.4 | 0.38 | 0.39 | 0.43 | 0.4003 | 0.4 |  | #N/A
| 1 | 3 | 0.36 | 0.425 | 0.4 | < 0.01 | 0.39545 | 0.4 | 0.4 | 0.42 | 0.4 | 0.4 | 0.4 | 0.42 | 0.417 | 0.39 |  | 16
| 1 | 4 | 0.36 | 0.429 | 0.42 | < 0.01 | 0.40014 | 0.41 | 0.4 | 0.42 | 0.4 | 0.38 | 0.39 | 0.41 | 0.4003 | 0.42 |  | 0.40653
| 1 | 5 | 0.36 | 0.429 | 0.41 | < 0.01 | 0.39942 | 0.41 | 0.4 | 0.41 | 0.41 | 0.39 | 0.38 | 0.44 | 0.417 | 0.42 |  | 133
| 1 | 6 | 0.37 | 0.428 | 0.41 | 0.58 | 0.40166 | 0.4 | 0.42 | 0.42 | 0.4 | 0.38 | 0.38 | 0.41 | 0.4003 | 0.4
| Mean |  | 0.36333 | 0.428 | 0.41167 | 0.58 | 0.40167 | 0.40667 | 0.405 | 0.41833 | 0.40167 | 0.38667 | 0.385 | 0.42 | 0.40865 | 0.405
| Median |  | 0.36 | 0.4285 | 0.41 | 0.58 | 0.4009 | 0.41 | 0.4 | 0.42 | 0.4 | 0.385 | 0.385 | 0.415 | 0.40865 | 0.4
| Std Dev. |  | 0.005164 | 0.0015492 | 0.0075277 |  | 0.0047975 | 0.005164 | 0.0083666 | 0.0040825 | 0.0040825 | 0.008165 | 0.010488 | 0.012649 | 0.009147 | 0.012247
| Rel.Std.Dev. |  | 1.421% | 0.362% | 1.829% |  | 1.194% | 1.27% | 2.066% | 0.9759% | 1.016% | 2.112% | 2.724% | 3.012% | 2.238% | 3.024%
| PDM3 |  | -10.63% | 5.281% | 1.264% | 42.67% | -1.196% | 0.03376% | -0.3762% | 2.904% | -1.196% | -4.886% | -5.296% | 3.314% | 0.5216% | -0.3762%
| Z-Score (Absolute) |  | 2.81 | 1.55 | 0.45 | 20.51 | 0.22 | 0.11 | 0 | 0.9 | 0.22 | 1.24 | 1.35 | 1.01 | 0.25 | 0
| Analytical results for Tl in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.25 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [13 empty] | 1
| 1 | 2 | < 100 | < 10 | [13 empty] | 15
| 1 | 3 | < 100 | < 10 | [13 empty] | 16
| 1 | 4 | < 100 | < 10 | [13 empty] | < 10
| 1 | 5 | < 100 | < 10 | [13 empty] | 21
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Tm in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Tm | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 8
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 22
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 313 (Indicative Value 12.8 ppm) | [16 empty] | Indicative
U | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | 13 | [14 empty] | 1
| 1 | 2 | 14 | [14 empty] | 17
| 1 | 3 | 12 | [14 empty] | 16
| 1 | 4 | 13 | [14 empty] | 12.8333
| 1 | 5 | 12 | [14 empty] | 23
| 1 | 6 | 13
| Mean |  | 12.8333
| Median |  | 13
| Std Dev. |  | 0.75277
| Rel.Std.Dev. |  | 5.866%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 313 (Indicative Value 127 ppm) | [16 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | 126.0453 | 100 | 200 | 100 | 100 | 151 | 73 | 166 | [6 empty] | 1
| 1 | 2 | < 100 | 125.9833 | 100 | 100 | 100 | 100 | 151 | 45 | 167 | [6 empty] | 18
| 1 | 3 | < 100 | 136.4 | 100 | 200 | 100 | 100 | 157 | 34 | 167 | [6 empty] | 16
| 1 | 4 | < 100 | 131.137 | 100 | 200 | 100 | 100 | 157 | 123 | 166 | [6 empty] | 127.2209
| 1 | 5 | < 100 | 135.543 | 200 | 100 | 100 | 100 | 146 | 112 | 168 | [6 empty] | 24
| 1 | 6 | < 100 | 132.287 | 100 | 200 | 100 | 100 | 157 | 112 | 168
| Mean |  |  | 131.2326 | 116.6667 | 166.6667 | 100 | 100 | 153.1667 | 83.1667 | 167
| Median |  |  | 131.712 | 100 | 200 | 100 | 100 | 154 | 92.5 | 167
| Std Dev. |  |  | 4.4915 | 40.8248 | 51.6398 | 0 | 0 | 4.5789 | 38.018 | 0.89443
| Rel.Std.Dev. |  |  | 3.423% | 34.99% | 30.98% | 0% | 0% | 2.99% | 45.71% | 0.5356%
| PDM3 |  |  | 3.153% | -8.296% | 31.01% | -21.4% | -21.4% | 20.39% | -34.63% | 31.27%
| Z-Score (Absolute) |  | 1.34 | 0.29 | 0 | 1 | 0.33 | 0.33 | 0.73 | 0.67 | 1.01
| Analytical results for W in OREAS 313 (Indicative Value 68 ppm) | [16 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | 30 | < 10 | < 10 | < 10 | 164 | [9 empty] | 1
| 1 | 2 | < 100 | 30 | < 10 | < 10 | < 10 | 162 | [9 empty] | 19
| 1 | 3 | < 100 | 30 | < 10 | < 10 | 10 | 157 | [9 empty] | 16
| 1 | 4 | < 100 | 40 | < 10 | < 10 | 20 | 162 | [9 empty] | 67.7222
| 1 | 5 | < 100 | 30 | < 10 | < 10 | 20 | 159 | [9 empty] | 25
| 1 | 6 | < 100 | 30 | < 10 | < 10 | < 10 | 160
| Mean |  |  | 31.6667 |  |  | 16.6667 | 160.6667
| Median |  |  | 30 |  |  | 20 | 161
| Std Dev. |  |  | 4.0825 |  |  | 5.7735 | 2.5033
| Rel.Std.Dev. |  |  | 12.89% |  |  | 34.64% | 1.558%
| PDM3 |  |  | -53.24% |  |  | -75.39% | 137.2%
| Z-Score (Absolute) |  | 1.19 | 0.43 | 0.67 | 0.67 | 0.43 | 5.79
| Analytical results for Y in OREAS 313 (Indicative Value 55 ppm) | [16 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 55 | 63 | 55 | < 39 | 71 | [10 empty] | 1
| 1 | 2 | 63 | < 39 | 47 | < 39 | 70 | [10 empty] | 20
| 1 | 3 | 63 | < 39 | 63 | < 39 | 70 | [10 empty] | 16
| 1 | 4 | 55 | 71 | 47 | < 39 | 70 | [10 empty] | 54.9879
| 1 | 5 | 63 | 55 | < 39 | < 39 | 71 | [10 empty] | 26
| 1 | 6 | 55 | 63 | < 39 | < 39 | 70
| Mean |  | 59 | 63 | 53 |  | 70.3333
| Median |  | 59 | 63 | 51 |  | 70
| Std Dev. |  | 4.3818 | 6.532 | 7.6594 |  | 0.5164
| Rel.Std.Dev. |  | 7.427% | 10.37% | 14.45% |  | 0.7342%
| PDM3 |  | 7.296% | 14.57% | -3.615% |  | 27.91%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.42 | 1.85 | 1.4
| Analytical results for Yb in OREAS 313 (Indicative Value < 10 ppm) | [16 empty] | Indicative
Yb | Round | Replicate | Lab | [14 empty] | 1
| No | No | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 10 | [14 empty] | 1
| 1 | 2 | < 10 | [14 empty] | 21
| 1 | 3 | < 10 | [14 empty] | 16
| 1 | 4 | < 10 | [14 empty] | < 10
| 1 | 5 | < 10 | [14 empty] | 27
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 313 (Certified Value 3.47 wt.%) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  |  | 2
3
| 1 | 1 | 3.5644 | 3.62 | 3.3841 | 3.41 | 3.53 | 3.44 | 3.43 | 3.76 | 3.49 | 3.48 | 3.54 | 3.2979 |  |  |  | 1
| 1 | 2 | 3.562 | 3.68 | 3.399 | 3.5 | 3.46 | 3.43 | 3.41 | 3.76 | 3.52 | 3.44 | 3.54 | 3.3047 |  |  |  | #N/A
| 1 | 3 | 3.5853 | 3.61 | 3.3736 | 3.52 | 3.46 | 3.43 | 3.46 | 3.68 | 3.48 | 3.46 | 3.54 | 3.3038 |  |  |  | 16
| 1 | 4 | 3.5516 | 3.48 | 3.3574 | 3.47 | 3.48 | 3.44 | 3.44 | 3.79 | 3.49 | 3.47 | 3.55 | 3.2945 |  |  |  | 3.4686
| 1 | 5 | 3.5572 | 3.56 | 3.3752 | 3.51 | 3.43 | 3.44 | 3.47 | 3.73 | 3.52 | 3.47 | 3.52 | 3.3028 |  |  |  | 134
| 1 | 6 | 3.5636 | 3.62 | 3.3221 | 3.48 | 3.56 | 3.39 | 3.43 | 3.74 | 3.5 | 3.45 | 3.5 | 3.277
| Mean |  | 3.564 | 3.595 | 3.3686 | 3.4817 | 3.4867 | 3.4283 | 3.44 | 3.7433 | 3.5 | 3.4617 | 3.5317 | 3.2968
| Median |  | 3.5628 | 3.615 | 3.3744 | 3.49 | 3.47 | 3.435 | 3.435 | 3.75 | 3.495 | 3.465 | 3.54 | 3.3004
| Std Dev. |  | 0.011472 | 0.068044 | 0.026526 | 0.039707 | 0.048854 | 0.019408 | 0.021909 | 0.037238 | 0.016733 | 0.01472 | 0.018348 | 0.010449
| Rel.Std.Dev. |  | 0.3219% | 1.893% | 0.7874% | 1.14% | 1.401% | 0.5661% | 0.6369% | 0.9948% | 0.4781% | 0.4252% | 0.5195% | 0.317%
| PDM3 |  | 2.751% | 3.645% | -2.884% | 0.3773% | 0.5215% | -1.16% | -0.8239% | 7.921% | 0.9059% | -0.1993% | 1.819% | -4.953%
| Z-Score (Absolute) |  | 1.04 | 1.45 | 1.51 | 0.03 | 0.03 | 0.73 | 0.58 | 3.38 | 0.21 | 0.29 | 0.62 | 2.45
| Analytical results for Zr in OREAS 313 (Indicative Value 106 ppm) | [16 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 185 | 200 | 100 | 100 | 100 | 194 | [8 empty] | 1
| 1 | 2 | 100 | < 185 | 200 | 100 | 100 | 100 | 194 | [8 empty] | 22
| 1 | 3 | < 100 | < 185 | 200 | 100 | 100 | 100 | 196 | [8 empty] | 16
| 1 | 4 | 200 | < 185 | 200 | 100 | 100 | 100 | 194 | [8 empty] | 106.25
| 1 | 5 | 200 | < 185 | 200 | 200 | 100 | 100 | 195 | [8 empty] | 28
| 1 | 6 | < 100 | < 185 | 200 | 100 | 100 | 100 | 192
| Mean |  | 166.6667 |  | 200 | 116.6667 | 100 | 100 | 194.1667
| Median |  | 200 |  | 200 | 100 | 100 | 100 | 194
| Std Dev. |  | 57.735 |  | 0 | 40.8248 | 0 | 0 | 1.3292
| Rel.Std.Dev. |  | 34.64% |  | 0% | 34.99% | 0% | 0% | 0.6845%
| PDM3 |  | 56.86% |  | 88.24% | 9.804% | -5.882% | -5.882% | 82.75%
| Z-Score (Absolute) |  | 0 | 1.28 | 7.42 | 0.67 | 0.67 | 0.67 | 6.95
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 313 (Certified Value 7.94 wt.%) | [18 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | OxBF*XRF | LOI*TGA | LOI*Furnace | BF*XRF | 2
3
| 1 | 1 | 7.9755 | 7.87 | 6.7 | 6.75 | 6.5555 | 7.98 | 8.04 | 7.73 | 7.25 | 8.16 | 8.0552 | 6.75 | 7.39 | 8.11 | 8.28 | 8.12 | 7.46 | 1
| 1 | 2 | 8.0046 | 7.83 | 6.68 | 7.38 | 6.6039 | 7.84 | 7.9 | 8.22 | 7.64 | 8.15 | 8.0984 | 6.79 | 7.51 | 8.2 | 8.3 | 8.1 | 7.47 | #N/A
| 1 | 3 | 8.0292 | 7.85 | 6.64 | 7.47 | 6.6629 | 7.82 | 8.09 | 8.13 | 7.46 | 8.26 | 8.0176 | 6.67 | 8.06 | 8.08 | 8.25 | 8.19 | 7.45 | 16
| 1 | 4 | 7.9541 | 7.83 | 6.52 | 6.82 | 6.6756 | 7.77 | 7.9 | 7.76 | 7.34 | 8.23 | 8.1594 | 6.82 | 8.37 | 8.07 | 8.21 | 8.02 | 7.49 | 7.9435
| 1 | 5 | 7.9682 | 7.75 | 6.78 | 7.6 | 6.5276 | 7.96 | 7.89 | 7.99 | 7.83 | 8.15 | 7.9376 | 6.72 | 8.17 | 8.13 | 8.3 | 8.18 | 7.47 | 136
| 1 | 6 | 8.003 | 7.75 | 6.74 | 6.89 | 6.606 | 7.89 | 7.94 | 8.1 | 7.44 | 8.17 | 8.1984 | 6.7 | 8.18 | 7.96 | 8.11 | 8.16 | 7.49
| Mean |  | 7.9891 | 7.8133 | 6.6767 | 7.1517 | 6.6053 | 7.8767 | 7.96 | 7.9883 | 7.4933 | 8.1867 | 8.0778 | 6.7417 | 7.9467 | 8.0917 | 8.2417 | 8.1283 | 7.4717
| Median |  | 7.9892 | 7.83 | 6.69 | 7.135 | 6.605 | 7.865 | 7.92 | 8.045 | 7.45 | 8.165 | 8.0768 | 6.735 | 8.115 | 8.095 | 8.265 | 8.14 | 7.47
| Std Dev. |  | 0.027872 | 0.051251 | 0.090701 | 0.37263 | 0.057924 | 0.082138 | 0.084617 | 0.20253 | 0.21049 | 0.046762 | 0.095292 | 0.056362 | 0.39923 | 0.079352 | 0.073052 | 0.063377 | 0.016021
| Rel.Std.Dev. |  | 0.3489% | 0.6559% | 1.358% | 5.21% | 0.8769% | 1.043% | 1.063% | 2.535% | 2.809% | 0.5712% | 1.18% | 0.836% | 5.024% | 0.9807% | 0.8864% | 0.7797% | 0.2144%
| PDM3 |  | 0.5744% | -1.638% | -15.95% | -9.968% | -16.85% | -0.8411% | 0.208% | 0.5647% | -5.667% | 3.062% | 1.69% | -15.13% | 0.04017% | 1.866% | 3.754% | 2.327% | -5.94%
| Z-Score (Absolute) |  | 0.16 | 0.49 | 4.71 | 2.95 | 4.98 | 0.26 | 0.05 | 0.15 | 1.68 | 0.89 | 0.49 | 4.47 | 0 | 0.54 | 1.09 | 0.67 | 1.76
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 313 (Indicative Value 76.8 ppm) | [22 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.04 | 2.13 | 2.23 | 2.24 | [17 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES | PF*OES/MS | [17 empty] | 1
1
| 1 | 1 | 76.89 | 76 | 77.33 | 75 | [17 empty] | 1
| 1 | 2 | 76.39 | 77 | 80.87 | 76 | [17 empty] | 10
| 1 | 3 | 77.34 | 77 | 79.68 | 76 | [17 empty] | 16
| 1 | 4 | 76.53 | 75 | 75.56 | 74 | [17 empty] | 76.7815
| 1 | 5 | 77.08 | 76 | 85.3 | 78 | [17 empty] | 32
| 1 | 6 | 75.77 | 78 | 79.69 | 73
| Mean |  | 76.6667 | 76.5 | 79.7383 | 75.3333
| Median |  | 76.71 | 76.5 | 79.685 | 75.5
| Std Dev. |  | 0.56095 | 1.0488 | 3.331 | 1.7512
| Rel.Std.Dev. |  | 0.7317% | 1.371% | 4.177% | 2.325%
| PDM3 |  | -0.1496% | -0.3666% | 3.851% | -1.886%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 3.19 | 1.26
| Analytical results for Al in OREAS 313 (Certified Value 4.96 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 4.73 | 4.8673 | 4.4 | 5.04 | 4.8753 | 4.89 | 4.768 | 4.98 | 4.933 | 4.848 | 4.97 | 4.88 | 5.0128 | 4.96 | 5.11 | 5.049 | 5.02 | 4.821 | 5.03 | 4.68 |  | 1
| 1 | 2 | 4.77 | 4.8486 | 4.44 | 5.07 | 4.8657 | 4.848 | 4.991 | 5.012 | 4.911 | 4.959 | 4.99 | 4.89 | 5.0064 | 4.96 | 5.12 | 5.007 | 5.03 | 4.808 | 4.97 | 4.66 |  | #N/A
| 1 | 3 | 4.73 | 4.8512 | 4.51 | 5.06 | 4.8155 | 4.938 | 5.038 | 4.901 | 4.959 | 5.001 | 5 | 4.91 | 5.0266 | 4.99 | 5.06 | 4.975 | 5 | 4.833 | 5 | 4.67 |  | 16
| 1 | 4 | 4.69 | 4.8205 | 4.52 | 5.07 | 4.8629 | 4.874 | 4.938 | 5.001 | 4.933 | 4.858 | 5.01 | 4.84 | 4.9795 | 5 | 5.17 | 4.948 | 5 | 4.834 | 4.97 | 4.64 |  | 4.9554
| 1 | 5 | 4.7 | 4.8637 | 4.44 | 5.03 | 4.9326 | 4.933 | 4.938 | 4.911 | 4.911 | 5.049 | 4.97 | 4.92 | 4.9925 | 4.95 | 5.06 | 5.054 | 4.91 | 4.818 | 4.96 | 4.84 |  | 138
| 1 | 6 | 4.72 | 4.8281 | 4.44 | 5.1 | 4.8497 | 4.895 | 5.001 | 4.98 | 4.901 | 4.911 | 4.97 | 4.91 | 5.0348 | 5 | 5.06 | 5.028 | 4.95 | 4.812 | 4.99 | 4.9
| Mean |  | 4.7233 | 4.8466 | 4.4583 | 5.0617 | 4.867 | 4.8963 | 4.9457 | 4.9642 | 4.9247 | 4.9377 | 4.985 | 4.8917 | 5.0087 | 4.9767 | 5.0967 | 5.0102 | 4.985 | 4.821 | 4.9867 | 4.7317
| Median |  | 4.725 | 4.8499 | 4.44 | 5.065 | 4.8643 | 4.8925 | 4.9645 | 4.98 | 4.922 | 4.935 | 4.98 | 4.9 | 5.0096 | 4.975 | 5.085 | 5.0175 | 5 | 4.8195 | 4.98 | 4.675
| Std Dev. |  | 0.028048 | 0.018798 | 0.046655 | 0.024833 | 0.038332 | 0.034518 | 0.095211 | 0.046825 | 0.021219 | 0.079949 | 0.017607 | 0.029269 | 0.020679 | 0.022509 | 0.045019 | 0.042092 | 0.045935 | 0.010696 | 0.02582 | 0.10962
| Rel.Std.Dev. |  | 0.5938% | 0.3879% | 1.046% | 0.4906% | 0.7876% | 0.705% | 1.925% | 0.9433% | 0.4309% | 1.619% | 0.3532% | 0.5983% | 0.4129% | 0.4523% | 0.8833% | 0.8401% | 0.9215% | 0.2219% | 0.5178% | 2.317%
| PDM3 |  | -4.682% | -2.195% | -10.03% | 2.145% | -1.784% | -1.191% | -0.1954% | 0.1779% | -0.6192% | -0.3569% | 0.5983% | -1.285% | 1.077% | 0.4301% | 2.852% | 1.106% | 0.5983% | -2.711% | 0.6319% | -4.514%
| Z-Score (Absolute) |  | 2.52 | 1.1 | 5.57 | 1.38 | 0.86 | 0.52 | 0.05 | 0.26 | 0.2 | 0.05 | 0.5 | 0.58 | 0.77 | 0.4 | 1.78 | 0.79 | 0.5 | 1.39 | 0.52 | 2.42
| Analytical results for As in OREAS 313 (Certified Value 93 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES | PF*OES | PF*MS | PF*OES/MS |  |  |  |  | 0
1
| 1 | 1 | 90 | 87.32 | 60.8 | 100 | 100 | 100 | 100 | 100 | 94 | 87 | 94.5245 | 96 | < 100 | 96 | 97 | 97 | 93 |  |  |  |  | 1
| 1 | 2 | 90 | 84.18 | 48.6 | 200 | 100 | 100 | 100 | 100 | 93 | 87 | 90.722 | 96 | < 100 | 95 | 96 | 102 | 91 |  |  |  |  | 11
| 1 | 3 | 90 | 83.43 | 62.7 | 100 | 100 | 100 | 100 | 100 | 97 | 86 | 89.8019 | 101 | < 100 | 93 | 96 | 97 | 97 |  |  |  |  | 16
| 1 | 4 | 90 | 89.09 | 57.6 | 100 | 100 | 100 | 100 | 100 | 92 | 84 | 91.0006 | 101 | < 100 | 93 | 96 | 97 | 92 |  |  |  |  | 92.6456
| 1 | 5 | 90 | 86.48 | 49.1 | 200 | 100 | 100 | 100 | < 100 | 98 | 86 | 93.9445 | 95 | 100 | 92 | 96 | 99 | 99 |  |  |  |  | 139
| 1 | 6 | 90 | 82.01 | 63.1 | 100 | 100 | 100 | 100 | 100 | 90 | 88 | 92.8311 | 97 | < 100 | 93 | 97 | 97 | 89
| Mean |  | 90 | 85.4183 | 56.9833 | 133.3333 | 100 | 100 | 100 | 100 | 94 | 86.3333 | 92.1374 | 97.6667 | 100 | 93.6667 | 96.3333 | 98.1667 | 93.5
| Median |  | 90 | 85.33 | 59.2 | 100 | 100 | 100 | 100 | 100 | 93.5 | 86.5 | 91.9159 | 96.5 | 100 | 93 | 96 | 97 | 92.5
| Std Dev. |  | 0 | 2.6581 | 6.595 | 51.6398 | 0 | 0 | 0 | 0 | 3.0332 | 1.3663 | 1.9076 | 2.6583 |  | 1.5055 | 0.5164 | 2.0412 | 3.7815
| Rel.Std.Dev. |  | 0% | 3.112% | 11.57% | 38.73% | 0% | 0% | 0% | 0% | 3.227% | 1.583% | 2.07% | 2.722% |  | 1.607% | 0.5361% | 2.079% | 4.044%
| PDM3 |  | -2.856% | -7.801% | -38.49% | 43.92% | 7.938% | 7.938% | 7.938% | 7.938% | 1.462% | -6.813% | -0.5485% | 5.42% | 7.938% | 1.102% | 3.98% | 5.959% | 0.9222%
| Z-Score (Absolute) |  | 0.67 | 1.56 | 7.04 | NA | NA | NA | NA | NA | 0.1 | 1.38 | 0.26 | 0.8 | NA | 0.03 | 0.55 | 0.9 | 0
Results from laboratories 2.07, 2.08, 2.09, 2.10, 2.11 and 2.17 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 313 (Indicative Value 96 ppm) | [22 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.12 | 2.14 | 2.16 | 2.18 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*MS | [17 empty] | 0
0
| 1 | 1 | 90 | 112.788 | 107 | 82 | [17 empty] | 1
| 1 | 2 | 120 | 111.346 | 98 | 99 | [17 empty] | 1
| 1 | 3 | 80 | 102.3453 | 102 | 91 | [17 empty] | 16
| 1 | 4 | 90 | 102.8253 | 106 | 93 | [17 empty] | 96.1422
| 1 | 5 | 80 | 109.0498 | 101 | 83 | [17 empty] | 33
| 1 | 6 | 80 | 103.0589 | 103 | 97
| Mean |  | 90 | 106.9022 | 102.8333 | 90.8333
| Median |  | 85 | 106.0544 | 102.5 | 92
| Std Dev. |  | 15.4919 | 4.7151 | 3.3116 | 7.0545
| Rel.Std.Dev. |  | 17.21% | 4.411% | 3.22% | 7.766%
| PDM3 |  | -6.389% | 11.19% | 6.96% | -5.522%
| Z-Score (Absolute) |  | 0.72 | 1.06 | 0.63 | 0.63
| Analytical results for Ba in OREAS 313 (Certified Value 0.238 wt.%) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [6 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*MS | PF*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.231 | 0.35611 | 0.2278 | 0.22764 | 0.23866 | 0.234 | 0.236 | 0.23876 | 0.2419 | 0.24 | 0.2431 | 0.2444 | 0.2022 | 0.2428 | 0.2267 | [6 empty] | 1
| 1 | 2 | 0.231 | 0.35623 | 0.2285 | 0.23413 | 0.23633 | 0.245 | 0.2357 | 0.24473 | 0.2481 | 0.24 | 0.245 | 0.2436 | 0.2037 | 0.2434 | 0.2247 | [6 empty] | 2
| 1 | 3 | 0.231 | 0.35493 | 0.2282 | 0.23331 | 0.23811 | 0.244 | 0.2369 | 0.24172 | 0.2461 | 0.24 | 0.2428 | 0.2494 | 0.2077 | 0.2403 | 0.2263 | [6 empty] | 16
| 1 | 4 | 0.234 | 0.35892 | 0.2262 | 0.23913 | 0.23872 | 0.245 | 0.2338 | 0.24003 | 0.2441 | 0.24 | 0.2425 | 0.2408 | 0.2078 | 0.2393 | 0.226 | [6 empty] | 0.23783
| 1 | 5 | 0.232 | 0.35579 | 0.2277 | 0.2329 | 0.24009 | 0.243 | 0.2379 | 0.24532 | 0.2457 | 0.24 | 0.2433 | 0.2442 | 0.2055 | 0.243 | 0.2382 | [6 empty] | 140
| 1 | 6 | 0.231 | 0.35566 | 0.2216 | 0.23733 | 0.23862 | 0.234 | 0.238 | 0.24249 | 0.2406 | 0.24 | 0.2428 | 0.2448 | 0.2027 | 0.2412 | 0.2348
| Mean |  | 0.23167 | 0.35627 | 0.22667 | 0.23407 | 0.23842 | 0.24083 | 0.23638 | 0.24217 | 0.24442 | 0.24 | 0.24325 | 0.24453 | 0.20493 | 0.24167 | 0.22945
| Median |  | 0.231 | 0.35595 | 0.22775 | 0.23372 | 0.23864 | 0.2435 | 0.23645 | 0.2421 | 0.2449 | 0.24 | 0.24295 | 0.2443 | 0.2046 | 0.242 | 0.2265
| Std Dev. |  | 0.0012111 | 0.0013753 | 0.0026059 | 0.0039881 | 0.0012199 | 0.0053448 | 0.0015791 | 0.0025681 | 0.0027946 | 0 | 0.00090056 | 0.0027818 | 0.0024566 | 0.0016585 | 0.005606
| Rel.Std.Dev. |  | 0.5228% | 0.386% | 1.15% | 1.704% | 0.5117% | 2.219% | 0.668% | 1.06% | 1.143% | 0% | 0.3702% | 1.138% | 1.199% | 0.6863% | 2.443%
| PDM3 |  | -2.591% | 49.8% | -4.694% | -1.579% | 0.2489% | 1.263% | -0.6081% | 1.827% | 2.77% | 0.9126% | 2.279% | 2.819% | -13.83% | 1.613% | -3.523%
| Z-Score (Absolute) |  | 1.27 | 17.75 | 2.04 | 0.9 | 0.24 | 0.13 | 0.55 | 0.33 | 0.67 | 0 | 0.5 | 0.69 | 5.35 | 0.25 | 1.61
| Analytical results for Be in OREAS 313 (Certified Value 2.35 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 2
2
| 1 | 1 | 2 | 3.1 | 2 | 2.2374 | < 3 | < 5 | 2.528 | 2 | 2.8 | 2 | < 5 | < 5 | 2 | < 2 | [7 empty] | 1
| 1 | 2 | 2 | 3 | 2.3 | 2.3028 | < 3 | < 5 | 2.5345 | 3 | 2.4 | 2 | < 5 | < 5 | 3 | 3 | [7 empty] | 33
| 1 | 3 | 2 | 3.1 | 1.9 | 2.3326 | < 3 | < 5 | 2.2531 | 2 | 3 | 2 | < 5 | < 5 | 3 | 3 | [7 empty] | 16
| 1 | 4 | 2 | 3 | 2.1 | 2.2182 | < 3 | < 5 | 2.3022 | 2 | 1.2 | 2 | < 5 | < 5 | 2 | 2 | [7 empty] | 2.3463
| 1 | 5 | 2 | 3 | 2 | 2.3302 | 3 | < 5 | 2.5819 | 3 | 3.6 | 2 | < 5 | < 5 | 3 | < 2 | [7 empty] | 141
| 1 | 6 | 2 | 3 | 2.3 | 2.2606 | < 3 | < 5 | 2.4842 | 2 | 1.8 | 2 | < 5 | < 5 | 2 | 3
| Mean |  | 2 | 3.0333 | 2.1 | 2.2803 | 3 |  | 2.4473 | 2.3333 | 2.4667 | 2 |  |  | 2.5 | 2.75
| Median |  | 2 | 3 | 2.05 | 2.2817 | 3 |  | 2.5061 | 2 | 2.6 | 2 |  |  | 2.5 | 3
| Std Dev. |  | 0 | 0.05164 | 0.16733 | 0.048618 |  |  | 0.13592 | 0.5164 | 0.8641 | 0 |  |  | 0.54772 | 0.5
| Rel.Std.Dev. |  | 0% | 1.702% | 7.968% | 2.132% |  |  | 5.554% | 22.13% | 35.03% | 0% |  |  | 21.91% | 18.18%
| PDM3 |  | -14.76% | 29.28% | -10.5% | -2.812% | 27.86% |  | 4.308% | -0.5514% | 5.131% | -14.76% |  |  | 6.552% | 17.21%
| Z-Score (Absolute) |  | NA | 11.46 | 7.42 | 3.77 | 14.5 | 0.67 | 0.39 | NA | 0 | NA | 0.67 | 0.67 | NA | NA
Results from laboratories 2.01, 2.16, 2.19, 2.23 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 313 (Certified Value 0.45 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.04 | 2.12 | 2.13 | 2.14 | 2.16 | 2.23 | 2.24 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [13 empty] | 2
2
| 1 | 1 | 0.5 | < 5 | < 2 | 0.4 | 0.42887 | 0.4 | 0.5 | 0.4 | [13 empty] | 1
| 1 | 2 | 0.5 | < 5 | < 2 | 0.5 | 0.38575 | 0.5 | 0.5 | 0.5 | [13 empty] | 34
| 1 | 3 | 0.5 | < 5 | < 2 | 0.4 | 0.39052 | 0.5 | 0.5 | 0.4 | [13 empty] | 16
| 1 | 4 | 0.5 | < 5 | < 2 | 0.4 | 0.41894 | 0.5 | 0.5 | 0.4 | [13 empty] | 0.44995
| 1 | 5 | 0.5 | < 5 | < 2 | 0.4 | 0.34123 | 0.5 | 0.5 | 0.5 | [13 empty] | 142
| 1 | 6 | 0.5 | < 5 | < 2 | 0.4 | 0.33272 | 0.4 | 0.4 | 0.5
| Mean |  | 0.5 |  |  | 0.41667 | 0.38301 | 0.46667 | 0.48333 | 0.45
| Median |  | 0.5 |  |  | 0.4 | 0.38814 | 0.5 | 0.5 | 0.45
| Std Dev. |  | 0 |  |  | 0.040825 | 0.039327 | 0.05164 | 0.040825 | 0.054772
| Rel.Std.Dev. |  | 0% |  |  | 9.798% | 10.27% | 11.07% | 8.447% | 12.17%
| PDM3 |  | 11.12% |  |  | -7.396% | -14.88% | 3.716% | 7.42% | 0.01211%
| Z-Score (Absolute) |  | 0.4 | 32.77 | 8.5 | 0.94 | 1.49 | 0.13 | 0.13 | 0.4
| Analytical results for Ca in OREAS 313 (Certified Value 1.34 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 1.3 | 1.3427 | 1.27 | 1.37 | 1.2876 | 1.079 | 1.279 | 1.258 | 1.301 | 1.272 | 1.37 | 1.3 | 1.3993 | 1.4 | 1.45 | 1.394 | 1.32 | 1.424 | 1.6 | 1.3 |  | 1
| 1 | 2 | 1.3 | 1.3111 | 1.28 | 1.38 | 1.3692 | 1.079 | 1.315 | 1.115 | 1.201 | 1.286 | 1.44 | 1.3 | 1.4073 | 1.4 | 1.5 | 1.379 | 1.34 | 1.415 | 1.6 | 1.3 |  | #N/A
| 1 | 3 | 1.3 | 1.3417 | 1.31 | 1.38 | 1.2841 | 1.115 | 1.358 | 1.115 | 1.215 | 1.365 | 1.46 | 1.3 | 1.4114 | 1.4 | 1.47 | 1.372 | 1.38 | 1.416 | 1.6 | 1.3 |  | 16
| 1 | 4 | 1.3 | 1.3533 | 1.3 | 1.38 | 1.3745 | 1.158 | 1.229 | 1.315 | 1.244 | 1.236 | 1.43 | 1.4 | 1.4099 | 1.4 | 1.48 | 1.358 | 1.32 | 1.427 | 1.6 | 1.4 |  | 1.3448
| 1 | 5 | 1.3 | 1.3411 | 1.28 | 1.38 | 1.3034 | 1.079 | 1.301 | 1.201 | 1.201 | 1.315 | 1.46 | 1.4 | 1.3934 | 1.4 | 1.39 | 1.387 | 1.34 | 1.42 | 1.6 | 1.4 |  | 143
| 1 | 6 | 1.3 | 1.3237 | 1.29 | 1.39 | 1.3445 | 1.172 | 1.329 | 1.244 | 1.179 | 1.279 | 1.32 | 1.3 | 1.3928 | 1.4 | 1.49 | 1.387 | 1.35 | 1.416 | 1.6 | 1.4
| Mean |  | 1.3 | 1.3356 | 1.2883 | 1.38 | 1.3272 | 1.1137 | 1.3018 | 1.208 | 1.2235 | 1.2922 | 1.4133 | 1.3333 | 1.4024 | 1.4 | 1.4633 | 1.3795 | 1.3417 | 1.4197 | 1.6 | 1.35
| Median |  | 1.3 | 1.3414 | 1.285 | 1.38 | 1.324 | 1.097 | 1.308 | 1.2225 | 1.208 | 1.2825 | 1.435 | 1.3 | 1.4033 | 1.4 | 1.475 | 1.383 | 1.34 | 1.418 | 1.6 | 1.35
| Std Dev. |  | 0 | 0.015308 | 0.01472 | 0.0063246 | 0.040711 | 0.042368 | 0.044481 | 0.080737 | 0.043569 | 0.043815 | 0.056451 | 0.05164 | 0.0083053 | 2.4324e-16 | 0.039833 | 0.012973 | 0.022286 | 0.0049261 | 2.4324e-16 | 0.054772
| Rel.Std.Dev. |  | 0% | 1.146% | 1.143% | 0.4583% | 3.067% | 3.804% | 3.417% | 6.683% | 3.561% | 3.391% | 3.994% | 3.873% | 0.5922% | 1.737e-14% | 2.722% | 0.9404% | 1.661% | 0.347% | 1.52e-14% | 4.057%
| PDM3 |  | -3.33% | -0.6829% | -4.198% | 2.618% | -1.306% | -17.19% | -3.194% | -10.17% | -9.019% | -3.913% | 5.097% | -0.8517% | 4.281% | 4.106% | 8.815% | 2.581% | -0.232% | 5.568% | 18.98% | 0.3876%
| Z-Score (Absolute) |  | 0.54 | 0.04 | 0.7 | 0.58 | 0.16 | 3.14 | 0.51 | 1.83 | 1.61 | 0.65 | 1.04 | 0.07 | 0.89 | 0.86 | 1.74 | 0.57 | 0.04 | 1.13 | 3.65 | 0.16
| Analytical results for Cd in OREAS 313 (Certified Value 112 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.02 | 2.04 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | [10 empty] | 0
0
| 1 | 1 | 110.98 | 110.3 | 110 | 112 | 108.4361 | 114 | 110 | 114 | 113 | 112 | 112 | [10 empty] | 1
| 1 | 2 | 109.12 | 109.4 | 113 | 113 | 116.3759 | 116 | 120 | 115 | 112 | 115 | 111 | [10 empty] | 13
| 1 | 3 | 109.28 | 111.5 | 111 | 111 | 105.6603 | 116 | 110 | 112 | 113 | 117 | 115 | [10 empty] | 16
| 1 | 4 | 112.13 | 110.9 | 111 | 109 | 115.0697 | 115 | 110 | 113 | 112 | 111 | 108 | [10 empty] | 111.6397
| 1 | 5 | 111.21 | 110.3 | 109 | 111 | 105.1647 | 117 | 110 | 110 | 112 | 116 | 115 | [10 empty] | 144
| 1 | 6 | 113.26 | 107.7 | 105 | 112 | 114.7562 | 115 | 110 | 112 | 113 | 116 | 111
| Mean |  | 110.9967 | 110.0167 | 109.8333 | 111.3333 | 110.9105 | 115.5 | 111.6667 | 112.6667 | 112.5 | 114.5 | 112
| Median |  | 111.095 | 110.3 | 110.5 | 111.5 | 111.5961 | 115.5 | 110 | 112.5 | 112.5 | 115.5 | 111.5
| Std Dev. |  | 1.6076 | 1.3333 | 2.7142 | 1.3663 | 5.0727 | 1.0488 | 4.0825 | 1.7512 | 0.54772 | 2.429 | 2.6833
| Rel.Std.Dev. |  | 1.448% | 1.212% | 2.471% | 1.227% | 4.574% | 0.9081% | 3.656% | 1.554% | 0.4869% | 2.121% | 2.396%
| PDM3 |  | -0.576% | -1.454% | -1.618% | -0.2744% | -0.6532% | 3.458% | 0.02417% | 0.9199% | 0.7706% | 2.562% | 0.3227%
| Z-Score (Absolute) |  | 0.49 | 1.21 | 1.35 | 0.25 | 0.56 | 2.82 | NA | 0.74 | 0.61 | 2.08 | 0.25
Results from laboratory 2.17 were removed due to their 10 ppm reading resolution.
| Analytical results for Ce in OREAS 313 (Certified Value 60 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 0
1
| 1 | 1 | 61.1 | 58.3 | 56.6788 | 54.4 | 60.4 | 51 | 59.6 | 60.1 | 61.6 | [12 empty] | 1
| 1 | 2 | 61.9 | 62.3 | 58.4601 | 61.4 | 62.7 | 54 | 59.3 | 60.5 | 62.3 | [12 empty] | 14
| 1 | 3 | 63.2 | 59.3 | 55.7281 | 62.9 | 59.2 | 57 | 59.4 | 60.9 | 62.1 | [12 empty] | 16
| 1 | 4 | 62.6 | 62.7 | 58.1813 | 61 | 60.8 | 58 | 58.7 | 60.3 | 60.8 | [12 empty] | 60.4291
| 1 | 5 | 63.2 | 61.4 | 56.3699 | 64.6 | 59.8 | 54 | 59.5 | 60.3 | 63.9 | [12 empty] | 145
| 1 | 6 | 61.3 | 61.3 | 56.4193 | 50.9 | 60.7 | 50 | 60.1 | 60 | 62
| Mean |  | 62.2167 | 60.8833 | 56.9729 | 59.2 | 60.6 | 54 | 59.4333 | 60.35 | 62.1167
| Median |  | 62.25 | 61.35 | 56.549 | 61.2 | 60.55 | 54 | 59.45 | 60.3 | 62.05
| Std Dev. |  | 0.92394 | 1.7279 | 1.0934 | 5.3453 | 1.1916 | 3.1623 | 0.45461 | 0.32094 | 1.0226
| Rel.Std.Dev. |  | 1.485% | 2.838% | 1.919% | 9.029% | 1.966% | 5.856% | 0.7649% | 0.5318% | 1.646%
| PDM3 |  | 2.958% | 0.7517% | -5.719% | -2.034% | 0.2828% | -10.64% | -1.648% | -0.1309% | 2.793%
| Z-Score (Absolute) |  | 1.09 | 0.31 | 1.98 | 0.67 | 0.15 | 3.72 | 0.54 | 0 | 1.04
| Analytical results for Co in OREAS 313 (Certified Value 26.4 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS |  | 1
2
| 1 | 1 | 23 | 19.8406 | 28.4 | 26.6 | 24.9697 | 40 | < 20 | 20 | 30 | 20 | 27.2 | 26.4 | 28.0109 | 27 | 30 | 26.6 | 28 | 24 | 27 | 27 |  | 1
| 1 | 2 | 23 | 19.0549 | 28.2 | 26.3 | 24.8547 | 40 | < 20 | 20 | 30 | 30 | 28.6 | 26.2 | 28.9003 | 26 | 30 | 26.4 | 27 | 24 | 28 | 27 |  | 16
| 1 | 3 | 22 | 22.0523 | 28.8 | 25.9 | 25.3801 | 40 | < 20 | < 20 | 40 | < 20 | 25.2 | 26.3 | 29.7829 | 27 | 30 | 26.5 | 27 | 25 | 27 | 28 |  | 16
| 1 | 4 | 23 | 21.1772 | 28.6 | 26.1 | 26.0165 | 40 | 20 | 30 | 30 | 20 | 24.4 | 26.1 | 28.8328 | 27 | 30 | 26.3 | 28 | 24 | 26 | 27 |  | 26.4422
| 1 | 5 | 23 | 19.6054 | 27.8 | 25.8 | 25.9609 | 40 | 20 | 30 | 20 | 20 | 25.6 | 26.7 | 28.049 | 26 | 30 | 26.8 | 27 | 24 | 28 | 29 |  | 146
| 1 | 6 | 23 | 22.1475 | 27.8 | 27.3 | 24.4726 | 30 | < 20 | < 20 | 20 | 20 | 26.8 | 26.4 | 29.5639 | 27 | 30 | 26.6 | 27 | 24 | 27 | 28
| Mean |  | 22.8333 | 20.6463 | 28.2667 | 26.3333 | 25.2758 | 38.3333 | 20 | 25 | 28.3333 | 22 | 26.3 | 26.35 | 28.8566 | 26.6667 | 30 | 26.5333 | 27.3333 | 24.1667 | 27.1667 | 27.6667
| Median |  | 23 | 20.5089 | 28.3 | 26.2 | 25.1749 | 40 | 20 | 25 | 30 | 20 | 26.2 | 26.35 | 28.8666 | 27 | 30 | 26.55 | 27 | 24 | 27 | 27.5
| Std Dev. |  | 0.40825 | 1.325 | 0.41312 | 0.55377 | 0.6237 | 4.0825 | 0 | 5.7735 | 7.5277 | 4.4721 | 1.5271 | 0.20736 | 0.73869 | 0.5164 | 0 | 0.17512 | 0.5164 | 0.40825 | 0.75277 | 0.8165
| Rel.Std.Dev. |  | 1.788% | 6.417% | 1.462% | 2.103% | 2.468% | 10.65% | 0% | 23.09% | 26.57% | 20.33% | 5.806% | 0.787% | 2.56% | 1.936% | 0% | 0.66% | 1.889% | 1.689% | 2.771% | 2.951%
| PDM3 |  | -13.65% | -21.92% | 6.9% | -0.4118% | -4.411% | 44.97% | -24.36% | -5.454% | 7.152% | -16.8% | -0.5379% | -0.3488% | 9.131% | 0.8488% | 13.45% | 0.3445% | 3.37% | -8.606% | 2.74% | 4.631%
| Z-Score (Absolute) |  | 2.37 | 3.8 | 1.2 | 0.07 | 0.76 | NA | NA | NA | NA | NA | 0.09 | 0.06 | 1.58 | 0.15 | NA | 0.06 | 0.58 | 1.49 | 0.48 | 0.8
| Analytical results for Cr in OREAS 313 (Certified Value 79 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES |  |  |  |  | 0
1
| 1 | 1 | 90 | 82 | 83.2991 | 70 | 140 | 270 | 70 | 140 | 80 | 75 | 71.7579 | 91 | 100 | 82 | 82 | 81 | 103 |  |  |  |  | 1
| 1 | 2 | 90 | 85 | 82.6163 | 70 | 140 | 140 | 70 | 140 | 90 | 83 | 72.8107 | 84 | 100 | 82 | 83 | 81 | 96 |  |  |  |  | 37
| 1 | 3 | 90 | 82 | 83.1236 | 70 | 70 | 140 | 70 | 140 | 80 | 83 | 73.6235 | 90 | 100 | 75 | 85 | 81 | 102 |  |  |  |  | 16
| 1 | 4 | 90 | 84 | 82.5962 | 70 | 140 | 140 | 70 | 140 | 80 | 81 | 76.0867 | 86 | 100 | 75 | 82 | 81 | 123 |  |  |  |  | 78.7981
| 1 | 5 | 90 | 82 | 82.0642 | 140 | 140 | 140 | 70 | 140 | 80 | 81 | 75.0572 | 82 | 100 | 82 | 85 | 81 | 123 |  |  |  |  | 147
| 1 | 6 | 90 | 82 | 82.7411 | 70 | 140 | 140 | 70 | 140 | 70 | 81 | 73.719 | 92 | 100 | 75 | 83 | 80 | 110
| Mean |  | 90 | 82.8333 | 82.7401 | 81.6667 | 128.3333 | 161.6667 | 70 | 140 | 80 | 80.6667 | 73.8425 | 87.5 | 100 | 78.5 | 83.3333 | 80.8333 | 109.5
| Median |  | 90 | 82 | 82.6787 | 70 | 140 | 140 | 70 | 140 | 80 | 81 | 73.6712 | 88 | 100 | 78.5 | 83 | 81 | 106.5
| Std Dev. |  | 0 | 1.3292 | 0.43645 | 28.5774 | 28.5774 | 53.0723 | 0 | 0 | 6.3246 | 2.9439 | 1.5483 | 4.0866 | 0 | 3.8341 | 1.3663 | 0.40825 | 11.3622
| Rel.Std.Dev. |  | 0% | 1.605% | 0.5275% | 34.99% | 22.27% | 32.83% | 0% | 0% | 7.906% | 3.649% | 2.097% | 4.67% | 0% | 4.884% | 1.64% | 0.505% | 10.38%
| PDM3 |  | 14.22% | 5.121% | 5.003% | 3.64% | 62.86% | 105.2% | -11.17% | 77.67% | 1.525% | 2.371% | -6.289% | 11.04% | 26.91% | -0.3783% | 5.755% | 2.583% | 38.96%
| Z-Score (Absolute) |  | NA | 0.01 | 0.01 | 0.45 | 6.53 | 11.76 | 2.2 | 8.27 | NA | 0.32 | 1.37 | 0.72 | NA | 0.63 | 0.08 | 0.3 | 4.09
Results from laboratories 2.01 and 2.12 were removed due to their 10 ppm reading resolution. Results from laboratory 2.17 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 313 (Certified Value 3.24 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 3.21 | 3.2 | 3.1737 | 3.5 | 3.2 | 2.8714 | 3.2 | 5 | 3.3 | 3.2 | 3.1 | [10 empty] | 1
| 1 | 2 | 3 | 3.3 | 3.1679 | 3.4 | 3.2 | 3.0165 | 3.3 | 23 | 3.3 | 3.3 | 3.1 | [10 empty] | 17
| 1 | 3 | 3.06 | 3.2 | 3.2011 | 3.6 | 3.3 | 3.0327 | 3.7 | 21 | 3.4 | 3.5 | 3.2 | [10 empty] | 16
| 1 | 4 | 3.14 | 3.3 | 3.3139 | 3.8 | 3.2 | 2.952 | 3.6 | 4 | 3.2 | 3 | 3.1 | [10 empty] | 3.2419
| 1 | 5 | 2.99 | 3.1 | 3.2907 | 3.2 | 3.6 | 2.8347 | 3.7 | 60 | 3.3 | 3.4 | 3.3 | [10 empty] | 148
| 1 | 6 | 3.07 | 3.3 | 3.2523 | 3.9 | 3.3 | 2.8629 | 3.2 | 13 | 3.3 | 3 | 3.2
| Mean |  | 3.0783 | 3.2333 | 3.2333 | 3.5667 | 3.3 | 2.9283 | 3.45 | 21 | 3.3 | 3.2333 | 3.1667
| Median |  | 3.065 | 3.25 | 3.2267 | 3.55 | 3.25 | 2.9117 | 3.45 | 17 | 3.3 | 3.25 | 3.15
| Std Dev. |  | 0.084242 | 0.08165 | 0.061661 | 0.2582 | 0.15492 | 0.084284 | 0.2429 | 20.6591 | 0.063246 | 0.20656 | 0.08165
| Rel.Std.Dev. |  | 2.737% | 2.525% | 1.907% | 7.239% | 4.695% | 2.878% | 7.041% | 98.38% | 1.917% | 6.388% | 2.578%
| PDM3 |  | -5.044% | -0.2632% | -0.2654% | 10.02% | 1.793% | -9.671% | 6.42% | 547.8% | 1.793% | -0.2632% | -2.32%
| Z-Score (Absolute) |  | 1.57 | 0 | 0 | 3.37 | 0.67 | 3.08 | 2.19 | 179.7 | 0.67 | 0 | 0.67
| Analytical results for Cu in OREAS 313 (Certified Value 0.036 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.036 | 0.041995 | 0.04129 | 0.044654 | 0.035 | 0.03 | 0.036 | 0.035 | 0.034 | 0.0349 | 0.035 | 0.036499 | 0.0359 | 0.04 | 0.0371 | 0.0346 | 0.0373 | 0.037 |  |  |  | 1
| 1 | 2 | 0.036 | 0.043149 | 0.0392 | 0.04469 | 0.035 | 0.032 | 0.037 | 0.034 | 0.035 | 0.0366 | 0.0346 | 0.037151 | 0.0358 | 0.04 | 0.0374 | 0.0345 | 0.0367 | 0.0365 |  |  |  | 39
| 1 | 3 | 0.036 | 0.042047 | 0.0388 | 0.044115 | 0.036 | 0.032 | 0.036 | 0.034 | 0.035 | 0.0378 | 0.0353 | 0.036424 | 0.0364 | 0.039 | 0.0355 | 0.0346 | 0.0368 | 0.037 |  |  |  | 16
| 1 | 4 | 0.035 | 0.04375 | 0.04055 | 0.045103 | 0.035 | 0.032 | 0.037 | 0.034 | 0.034 | 0.0356 | 0.0352 | 0.036439 | 0.0364 | 0.04 | 0.0363 | 0.0346 | 0.0372 | 0.0346 |  |  |  | 0.036316
| 1 | 5 | 0.035 | 0.043682 | 0.04385 | 0.043891 | 0.034 | 0.031 | 0.036 | 0.034 | 0.035 | 0.0381 | 0.0358 | 0.036904 | 0.0356 | 0.04 | 0.0354 | 0.0347 | 0.0373 | 0.0374 |  |  |  | 149
| 1 | 6 | 0.036 | 0.044407 | 0.0416 | 0.045637 | 0.034 | 0.032 | 0.037 | 0.034 | 0.034 | 0.0352 | 0.0355 | 0.036563 | 0.0364 | 0.039 | 0.0357 | 0.0347 | 0.0365 | 0.0369
| Mean |  | 0.035667 | 0.043172 | 0.040882 | 0.044682 | 0.034833 | 0.0315 | 0.0365 | 0.034167 | 0.0345 | 0.036367 | 0.035233 | 0.036663 | 0.036083 | 0.039667 | 0.036233 | 0.034617 | 0.036967 | 0.036567
| Median |  | 0.036 | 0.043416 | 0.04092 | 0.044672 | 0.035 | 0.032 | 0.0365 | 0.034 | 0.0345 | 0.0361 | 0.03525 | 0.036531 | 0.03615 | 0.04 | 0.036 | 0.0346 | 0.037 | 0.03695
| Std Dev. |  | 0.0005164 | 0.00097678 | 0.0018311 | 0.00063828 | 0.00075277 | 0.00083666 | 0.00054772 | 0.00040825 | 0.00054772 | 0.0013574 | 0.00041312 | 0.00029676 | 0.00036009 | 0.0005164 | 0.00085245 | 7.5277e-05 | 0.00034448 | 0.0010053
| Rel.Std.Dev. |  | 1.448% | 2.263% | 4.479% | 1.429% | 2.161% | 2.656% | 1.501% | 1.195% | 1.588% | 3.733% | 1.173% | 0.8094% | 0.9979% | 1.302% | 2.353% | 0.2175% | 0.9319% | 2.749%
| PDM3 |  | -1.789% | 18.88% | 12.57% | 23.03% | -4.083% | -13.26% | 0.5058% | -5.919% | -5.001% | 0.1387% | -2.982% | 0.9552% | -0.6415% | 9.226% | -0.2285% | -4.68% | 1.791% | 0.6894%
| Z-Score (Absolute) |  | 0.34 | 3.66 | 2.44 | 4.46 | 0.78 | 2.56 | 0.11 | 1.14 | 0.96 | 0.04 | 0.57 | 0.19 | 0.12 | 1.79 | 0.04 | 0.9 | 0.35 | 0.14
| Analytical results for Dy in OREAS 313 (Certified Value 3.57 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 3.5 | 3.7 | 3.5164 | 3.4 | 3.67 | 3 | 3.66 | 3.45 | 3.24 | [12 empty] | 1
| 1 | 2 | 3.7 | 3.7 | 3.9486 | 4.3 | 3.76 | 3 | 3.63 | 3.59 | 3.38 | [12 empty] | 5
| 1 | 3 | 3.6 | 3.8 | 3.1572 | 3.4 | 3.56 | 4 | 3.58 | 3.61 | 3.45 | [12 empty] | 16
| 1 | 4 | 3.6 | 4 | 3.9741 | 3.5 | 3.7 | 4 | 3.72 | 3.6 | 3.4 | [12 empty] | 3.573
| 1 | 5 | 3.3 | 3.6 | 3.718 | 3 | 3.73 | 3 | 3.5 | 3.61 | 3.59 | [12 empty] | 150
| 1 | 6 | 3.3 | 3.9 | 4.0873 | 3.2 | 3.67 | 3 | 3.59 | 3.6 | 3.31
| Mean |  | 3.5 | 3.7833 | 3.7336 | 3.4667 | 3.6817 | 3.3333 | 3.6133 | 3.5767 | 3.395
| Median |  | 3.55 | 3.75 | 3.8333 | 3.4 | 3.685 | 3 | 3.61 | 3.6 | 3.39
| Std Dev. |  | 0.16733 | 0.1472 | 0.34896 | 0.44572 | 0.069113 | 0.5164 | 0.075277 | 0.062503 | 0.12046
| Rel.Std.Dev. |  | 4.781% | 3.891% | 9.346% | 12.86% | 1.877% | 15.49% | 2.083% | 1.748% | 3.548%
| PDM3 |  | -2.042% | 5.888% | 4.497% | -2.975% | 3.043% | -6.706% | 1.13% | 0.104% | -4.981%
| Z-Score (Absolute) |  | 0.57 | 1.14 | 0.84 | 0.77 | 0.52 | NA | 0.11 | 0.11 | 1.21
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 313 (Certified Value 2.13 ppm) | [22 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 2 | 2.2 | 1.7854 | 2.2 | 2.05 | < 2 | 2.2 | 2.08 | 2.06 | [12 empty] | 1
| 1 | 2 | 2 | 2 | 1.8819 | 2.6 | 2.17 | 2 | 2.22 | 1.99 | 2.14 | [12 empty] | 6
| 1 | 3 | 2.1 | 2.3 | 1.8055 | 2.5 | 2.07 | 2 | 2.13 | 2.04 | 2.06 | [12 empty] | 16
| 1 | 4 | 2 | 2.1 | 1.8544 | 2.5 | 2.23 | 2 | 2.17 | 2.06 | 2.13 | [12 empty] | 2.125
| 1 | 5 | 1.9 | 2.2 | 1.8919 | 2 | 2.12 | 2 | 2.14 | 2.06 | 2.21 | [12 empty] | 151
| 1 | 6 | 1.9 | 2.2 | 1.869 | 2.1 | 2.16 | < 2 | 2.2 | 2.01 | 2.09
| Mean |  | 1.9833 | 2.1667 | 1.848 | 2.3167 | 2.1333 | 2 | 2.1767 | 2.04 | 2.115
| Median |  | 2 | 2.2 | 1.8617 | 2.35 | 2.14 | 2 | 2.185 | 2.05 | 2.11
| Std Dev. |  | 0.075277 | 0.10328 | 0.04308 | 0.24833 | 0.067132 | 0 | 0.036148 | 0.034059 | 0.057533
| Rel.Std.Dev. |  | 3.795% | 4.767% | 2.331% | 10.72% | 3.147% | 0% | 1.661% | 1.67% | 2.72%
| PDM3 |  | -6.667% | 1.961% | -13.03% | 9.02% | 0.3922% | -5.882% | 2.431% | -4% | -0.4706%
| Z-Score (Absolute) |  | 1.39 | 0.42 | 2.72 | 1.9 | 0.09 | NA | 0.52 | 0.83 | 0.09
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 313 (Certified Value 1 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.24 | [13 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | [13 empty] | 2
3
| 1 | 1 | 1.1 | 0.9 | 1.3453 | 1 | 0.96 | < 2 | 0.95 | 0.95 | [13 empty] | 1
| 1 | 2 | 1.1 | 1 | 1.3436 | 1.1 | 0.98 | < 2 | 0.97 | 0.91 | [13 empty] | 7
| 1 | 3 | 1.1 | 0.9 | 1.3144 | 1.1 | 0.93 | < 2 | 0.96 | 0.95 | [13 empty] | 16
| 1 | 4 | 1.1 | 0.9 | 1.3275 | 1.2 | 0.97 | < 2 | 0.93 | 0.94 | [13 empty] | 1.0036
| 1 | 5 | 1.1 | 1.1 | 1.3263 | 1.2 | 1.01 | < 2 | 0.92 | 1.02 | [13 empty] | 152
| 1 | 6 | 1.1 | 1.1 | 1.3095 | 0.9 | 0.97 | < 2 | 0.94 | 0.95
| Mean |  | 1.1 | 0.98333 | 1.3277 | 1.0833 | 0.97 |  | 0.945 | 0.95333
| Median |  | 1.1 | 0.95 | 1.3269 | 1.1 | 0.97 |  | 0.945 | 0.95
| Std Dev. |  | 2.4324e-16 | 0.098319 | 0.014649 | 0.1169 | 0.026077 |  | 0.018708 | 0.036148
| Rel.Std.Dev. |  | 2.211e-14% | 9.999% | 1.103% | 10.79% | 2.688% |  | 1.98% | 3.792%
| PDM3 |  | 9.604% | -2.02% | 32.3% | 7.944% | -3.349% |  | -5.84% | -5.01%
| Z-Score (Absolute) |  | 1.72 | 0.13 | 5.33 | 1.45 | 0.34 | 0.13 | 0.74 | 0.61
| Analytical results for Fe in OREAS 313 (Certified Value 6.23 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 6.17 | 6.3866 | 5.94 | 6.31 | 6.1546 | 6.099 | 5.987 | 6.218 | 6.099 | 5.987 | 6.38 | 5.93 | 6.2521 | 6.4 | 6.29 | 5.8939 | 6.344 | 6.39 | 6.514 | 6.22 | 5.94 | 1
| 1 | 2 | 6.18 | 6.3594 | 6.03 | 6.32 | 6.1591 | 5.987 | 6.281 | 6.148 | 6.043 | 6.197 | 6.39 | 6 | 6.3776 | 6.39 | 6.31 | 6.4096 | 6.309 | 6.3 | 6.509 | 6.14 | 5.92 | #N/A
| 1 | 3 | 6.17 | 6.36 | 6.18 | 6.33 | 6.0385 | 6.057 | 6.302 | 6.078 | 6.05 | 6.197 | 6.42 | 6.09 | 6.2457 | 6.4 | 6.18 | 6.6889 | 6.267 | 6.42 | 6.514 | 6.18 | 5.96 | 16
| 1 | 4 | 6.12 | 6.3795 | 6.18 | 6.33 | 6.1399 | 6.043 | 6.113 | 6.239 | 6.071 | 6.043 | 6.41 | 6.08 | 6.339 | 6.44 | 6.29 | 6.7107 | 6.323 | 6.29 | 6.521 | 6.14 | 5.97 | 6.2266
| 1 | 5 | 6.17 | 6.3498 | 6.03 | 6.31 | 6.1299 | 6.078 | 6.022 | 6.134 | 6.071 | 6.351 | 6.4 | 6.16 | 6.3307 | 6.35 | 6.26 | 6.3637 | 6.288 | 6.32 | 6.524 | 6.15 | 6.26 | 153
| 1 | 6 | 6.18 | 6.3799 | 5.99 | 6.38 | 6.1342 | 6.064 | 6.211 | 6.141 | 6.05 | 6.183 | 6.44 | 5.6 | 6.3905 | 6.43 | 6.32 | 5.9556 | 6.316 | 6.33 | 6.509 | 6.17 | 6.17
| Mean |  | 6.165 | 6.3692 | 6.0583 | 6.33 | 6.126 | 6.0547 | 6.1527 | 6.1597 | 6.064 | 6.1597 | 6.4067 | 5.9767 | 6.3226 | 6.4017 | 6.275 | 6.3371 | 6.3078 | 6.3417 | 6.5152 | 6.1667 | 6.0367
| Median |  | 6.17 | 6.3697 | 6.03 | 6.325 | 6.1371 | 6.0605 | 6.162 | 6.1445 | 6.0605 | 6.19 | 6.405 | 6.04 | 6.3348 | 6.4 | 6.29 | 6.3866 | 6.3125 | 6.325 | 6.514 | 6.16 | 5.965
| Std Dev. |  | 0.022583 | 0.014711 | 0.099883 | 0.026077 | 0.04436 | 0.038255 | 0.13286 | 0.059177 | 0.020765 | 0.12912 | 0.021602 | 0.20086 | 0.0614 | 0.031885 | 0.050892 | 0.34965 | 0.027081 | 0.051929 | 0.0061779 | 0.030768 | 0.14208
| Rel.Std.Dev. |  | 0.3663% | 0.231% | 1.649% | 0.412% | 0.7241% | 0.6318% | 2.159% | 0.9607% | 0.3424% | 2.096% | 0.3372% | 3.361% | 0.9711% | 0.4981% | 0.811% | 5.518% | 0.4293% | 0.8189% | 0.09482% | 0.4989% | 2.354%
| PDM3 |  | -0.9892% | 2.29% | -2.702% | 1.661% | -1.615% | -2.761% | -1.187% | -1.075% | -2.611% | -1.075% | 2.892% | -4.014% | 1.542% | 2.812% | 0.7775% | 1.774% | 1.305% | 1.848% | 4.635% | -0.9624% | -3.05%
| Z-Score (Absolute) |  | 0.01 | 1.05 | 0.56 | 0.85 | 0.21 | 0.58 | 0.07 | 0.04 | 0.53 | 0.04 | 1.24 | 0.99 | 0.81 | 1.22 | 0.56 | 0.88 | 0.73 | 0.91 | 1.81 | 0 | 0.67
| Analytical results for Ga in OREAS 313 (Certified Value 14.6 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.19 | 2.20 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 14 | 15 | 13.5 | 14.0824 | 16.3 | 15 | 15.1178 | 14 | 13.7 | 14 | 15 | 14 | [9 empty] | 1
| 1 | 2 | 14 | 15 | 14 | 14.1115 | 15.2 | 15 | 15.4957 | 15 | 13.7 | 16 | 15 | 14 | [9 empty] | 21
| 1 | 3 | 13 | 15 | 13.7 | 13.7968 | 16.6 | 15 | 14.9357 | 14 | 13.8 | 17 | 14 | 14 | [9 empty] | 16
| 1 | 4 | 14 | 14.9 | 14.9 | 14.3244 | 16.1 | 15 | 14.7373 | 14 | 13.8 | 16 | 15 | 14 | [9 empty] | 14.5973
| 1 | 5 | 14 | 14.7 | 14.5 | 13.9837 | 18.4 | 15 | 15.6913 | 14 | 13.7 | 16 | 15 | 15 | [9 empty] | 154
| 1 | 6 | 14 | 14.9 | 14 | 13.998 | 14.8 | 15 | 15.4343 | 14 | 13.7 | 15 | 13 | 14
| Mean |  | 13.8333 | 14.9167 | 14.1 | 14.0495 | 16.2333 | 15 | 15.2354 | 14.1667 | 13.7333 | 15.6667 | 14.5 | 14.1667
| Median |  | 14 | 14.95 | 14 | 14.0402 | 16.2 | 15 | 15.276 | 14 | 13.7 | 16 | 15 | 14
| Std Dev. |  | 0.40825 | 0.1169 | 0.51769 | 0.17396 | 1.2628 | 0 | 0.36523 | 0.40825 | 0.05164 | 1.0328 | 0.83666 | 0.40825
| Rel.Std.Dev. |  | 2.951% | 0.7837% | 3.672% | 1.238% | 7.779% | 0% | 2.397% | 2.882% | 0.376% | 6.592% | 5.77% | 2.882%
| PDM3 |  | -5.234% | 2.188% | -3.407% | -3.753% | 11.21% | 2.758% | 4.371% | -2.95% | -5.919% | 7.325% | -0.6669% | -2.95%
| Z-Score (Absolute) |  | 0.62 | 0.73 | 0.29 | 0.35 | 2.37 | 0.83 | 1.12 | 0.21 | 0.75 | 1.66 | 0.21 | 0.21
| Analytical results for Gd in OREAS 313 (Certified Value 4.37 ppm) | [22 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.24 | [13 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | [13 empty] | 2
3
| 1 | 1 | 4.5 | 4.03 | 3.3011 | 4.3 | 4.14 | 4 | 4.39 | 5 | [13 empty] | 1
| 1 | 2 | 4.5 | 4.39 | 3.5331 | 5 | 4.37 | 4 | 4.46 | 5 | [13 empty] | 8
| 1 | 3 | 4.5 | 4.29 | 3.452 | 4 | 4.14 | 4 | 4.24 | 5 | [13 empty] | 16
| 1 | 4 | 4.6 | 4.36 | 3.53 | 4.7 | 4.32 | 5 | 4.16 | 5 | [13 empty] | 4.3683
| 1 | 5 | 4.4 | 4.24 | 3.5403 | 5.1 | 4.23 | 4 | 4.15 | 5 | [13 empty] | 155
| 1 | 6 | 4.4 | 4.14 | 3.5226 | 4.4 | 4.26 | 4 | 4.34 | 5
| Mean |  | 4.4833 | 4.2417 | 3.4798 | 4.5833 | 4.2433 | 4.1667 | 4.29 | 5
| Median |  | 4.5 | 4.265 | 3.5263 | 4.55 | 4.245 | 4 | 4.29 | 5
| Std Dev. |  | 0.075277 | 0.13674 | 0.093321 | 0.42622 | 0.093524 | 0.40825 | 0.12681 | 0
| Rel.Std.Dev. |  | 1.679% | 3.224% | 2.682% | 9.299% | 2.204% | 9.798% | 2.956% | 0%
| PDM3 |  | 2.633% | -2.9% | -20.34% | 4.922% | -2.862% | -4.617% | -1.793% | 14.46%
| Z-Score (Absolute) |  | 1.21 | 0.14 | 4.39 | 1.77 | 0.13 | NA | 0.13 | NA
Results from laboratories 2.18 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 313 (Certified Value 10.4 ppm) | [22 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.23 | 2.24 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [13 empty] | 1
2
| 1 | 1 | 11 | 10.1538 | 10.7 | 10 | 9.4426 | 11 | 10 | 12 | [13 empty] | 1
| 1 | 2 | 11 | 10.5173 | 10.9 | 10 | 10.1289 | 11 | 11 | 12 | [13 empty] | 23
| 1 | 3 | 10 | 10.3857 | 10.6 | 10 | 9.3164 | 11 | 11 | 12 | [13 empty] | 16
| 1 | 4 | 11 | 10.8938 | 10.2 | 10 | 9.7661 | 11 | 10 | 11 | [13 empty] | 10.4054
| 1 | 5 | 11 | 10.9064 | 10.3 | 10 | 9.9693 | 11 | 10 | 12 | [13 empty] | 156
| 1 | 6 | 11 | 10.4001 | 9.7 | 10 | 9.7467 | 10 | 11 | 12
| Mean |  | 10.8333 | 10.5428 | 10.4 | 10 | 9.7284 | 10.8333 | 10.5 | 11.8333
| Median |  | 11 | 10.4587 | 10.45 | 10 | 9.7564 | 11 | 10.5 | 12
| Std Dev. |  | 0.40825 | 0.30082 | 0.42895 | 0 | 0.30714 | 0.40825 | 0.54772 | 0.40825
| Rel.Std.Dev. |  | 3.768% | 2.853% | 4.125% | 0% | 3.157% | 3.768% | 5.216% | 3.45%
| PDM3 |  | 4.113% | 1.321% | -0.05199% | -3.896% | -6.507% | 4.113% | 0.9091% | 13.72%
| Z-Score (Absolute) |  | 0.67 | 0.05 | 0.26 | 1.13 | 1.71 | 0.67 | 0.05 | 2.84
| Analytical results for Hf in OREAS 313 (Certified Value 3.02 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.19 | 2.22 | [14 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | BF*MS | [14 empty] | 2
3
| 1 | 1 | 1.3 | 2.9 | 3.2232 | < 10 | 3 | 3.1 | 2.7 | [14 empty] | 1
| 1 | 2 | 1.4 | 2.9 | 3.339 | < 10 | 3 | 3 | 2.8 | [14 empty] | 30
| 1 | 3 | 1.3 | 3 | 3.2922 | < 10 | 3 | 3 | 2.8 | [14 empty] | 16
| 1 | 4 | 1.4 | 3.3 | 3.3571 | < 10 | 3 | 3 | 2.8 | [14 empty] | 3.022
| 1 | 5 | 1.4 | 3.1 | 3.4084 | < 10 | 3 | 3 | 2.7 | [14 empty] | 157
| 1 | 6 | 1.6 | 3 | 3.3412 | < 10 | 3 | 2.9 | 2.7
| Mean |  | 1.4 | 3.0333 | 3.3269 |  | 3 | 3 | 2.75
| Median |  | 1.4 | 3 | 3.3401 |  | 3 | 3 | 2.75
| Std Dev. |  | 0.10954 | 0.15055 | 0.063049 |  | 0 | 0.063246 | 0.054772
| Rel.Std.Dev. |  | 7.825% | 4.963% | 1.895% |  | 0% | 2.108% | 1.992%
| PDM3 |  | -53.67% | 0.3738% | 10.09% |  | -0.7292% | -0.7292% | -9.002%
| Z-Score (Absolute) |  | 4.32 | 0.09 | 0.88 | 5.39 | 0 | 0 | 0.67
| Analytical results for Hg in OREAS 313 (Indicative Value < 5 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | Lab | [20 empty] | 1
| No | No | 2.04 | [20 empty] | ppm
|  |  | PF*OES | [20 empty] | 2
2
| 1 | 1 | < 5 | [20 empty] | 1
| 1 | 2 | < 5 | [20 empty] | 24
| 1 | 3 | < 5 | [20 empty] | 16
| 1 | 4 | < 5 | [20 empty] | < 5
| 1 | 5 | < 5 | [20 empty] | 34
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 313 (Certified Value 0.72 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 0.7 | 0.7 | 0.69327 | 0.7 | 0.76 | < 2 | 0.7 | 0.71 | 0.7 | [12 empty] | 1
| 1 | 2 | 0.7 | 0.7 | 0.71147 | 0.7 | 0.76 | < 2 | 0.71 | 0.71 | 0.71 | [12 empty] | 4
| 1 | 3 | 0.7 | 0.7 | 0.68666 | 0.8 | 0.73 | < 2 | 0.7 | 0.74 | 0.69 | [12 empty] | 16
| 1 | 4 | 0.7 | 0.7 | 0.73011 | 0.8 | 0.74 | < 2 | 0.72 | 0.74 | 0.72 | [12 empty] | 0.71869
| 1 | 5 | 0.6 | 0.8 | 0.7138 | 0.7 | 0.75 | < 2 | 0.71 | 0.7 | 0.73 | [12 empty] | 158
| 1 | 6 | 0.6 | 0.7 | 0.69532 | 0.6 | 0.77 | < 2 | 0.71 | 0.71 | 0.71
| Mean |  | 0.66667 | 0.71667 | 0.70511 | 0.71667 | 0.75167 |  | 0.70833 | 0.71833 | 0.71
| Median |  | 0.7 | 0.7 | 0.70339 | 0.7 | 0.755 |  | 0.71 | 0.71 | 0.71
| Std Dev. |  | 0.05164 | 0.040825 | 0.016232 | 0.075277 | 0.01472 |  | 0.0075277 | 0.017224 | 0.014142
| Rel.Std.Dev. |  | 7.746% | 5.696% | 2.302% | 10.5% | 1.958% |  | 1.063% | 2.398% | 1.992%
| PDM3 |  | -7.238% | -0.2813% | -1.89% | -0.2813% | 4.589% |  | -1.441% | -0.04935% | -1.209%
| Z-Score (Absolute) |  | NA | NA | 0.82 | NA | 3.4 | 25.88 | 0.53 | 0.38 | 0.38
Results from laboratories 2.04, 2.05 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 313 (Certified Value 0.3 ppm) | [22 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.23 | 2.24 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [13 empty] | 2
3
| 1 | 1 | 0.3 | < 0.5 | 0.3 | 0.3 | 0.20349 | 0.3 | 0.3 | < 0.3 | [13 empty] | 1
| 1 | 2 | 0.3 | < 0.5 | 0.6 | 0.3 | 0.32991 | 0.3 | 0.3 | < 0.3 | [13 empty] | 11
| 1 | 3 | 0.3 | < 0.5 | 0.4 | 0.3 | 0.30022 | 0.3 | 0.3 | < 0.3 | [13 empty] | 16
| 1 | 4 | 0.3 | < 0.5 | 0.4 | 0.3 | 0.38466 | 0.3 | 0.4 | < 0.3 | [13 empty] | 0.30327
| 1 | 5 | 0.3 | < 0.5 | 0.4 | 0.3 | 0.31716 | 0.3 | 0.3 | < 0.3 | [13 empty] | 159
| 1 | 6 | 0.3 | < 0.5 | 0.3 | 0.3 | 0.2499 | 0.3 | 0.3 | < 0.3
| Mean |  | 0.3 |  | 0.4 | 0.3 | 0.29755 | 0.3 | 0.31667
| Median |  | 0.3 |  | 0.4 | 0.3 | 0.30869 | 0.3 | 0.3
| Std Dev. |  | 0 |  | 0.10954 | 0 | 0.063475 | 0 | 0.040825
| Rel.Std.Dev. |  | 0% |  | 27.39% | 0% | 21.33% | 0% | 12.89%
| PDM3 |  | -1.079% |  | 31.89% | -1.079% | -1.886% | -1.079% | 4.416%
| Z-Score (Absolute) |  | 0 | 3.53 | 7.06 | 0 | 0.17 | 0 | 1.18 | 10.58
| Analytical results for K in OREAS 313 (Certified Value 2.05 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 2 | 2.0329 | 1.77 | 2.04 | 2.0869 | 2.042 | 1.86 | 2.084 | 2.009 | 2.05 | 2.1 | 2 | 1.8957 | 2.02 | 2.1 | 1.945 | 2.034 | 2.01 | 2.035 | 2.11 | 1.99 | 1
| 1 | 2 | 2.09 | 2.0562 | 1.78 | 2.05 | 2.0415 | 2.059 | 1.918 | 2.1 | 2.009 | 2.017 | 2.1 | 2 | 1.9624 | 2.03 | 2 | 2.1224 | 2.05 | 2.03 | 2.038 | 2.08 | 1.99 | #N/A
| 1 | 3 | 2.1 | 2.0763 | 1.81 | 2.05 | 2.0467 | 2.092 | 1.901 | 2.009 | 2.017 | 2.05 | 2.1 | 2 | 1.9675 | 2.05 | 2.1 | 2.1669 | 2.05 | 2.07 | 2.035 | 2.1 | 1.99 | 16
| 1 | 4 | 2.05 | 2.0411 | 1.81 | 2.06 | 2.0597 | 2.059 | 1.826 | 2.067 | 2.009 | 1.976 | 2.1 | 2.1 | 1.9592 | 2.04 | 2.1 | 2.1194 | 2.067 | 1.99 | 2.036 | 2.09 | 1.98 | 2.0466
| 1 | 5 | 2.08 | 2.0353 | 1.78 | 2.04 | 2.0591 | 2.092 | 1.901 | 2.059 | 2.034 | 2.059 | 2.1 | 2.1 | 1.8922 | 2.01 | 2 | 2.1294 | 2.034 | 2.03 | 2.032 | 2.07 | 2.05 | 160
| 1 | 6 | 2.02 | 2.0551 | 1.79 | 2.07 | 2.0691 | 2.067 | 1.901 | 2.042 | 2.009 | 1.959 | 2.1 | 2 | 1.9691 | 2.04 | 2 | 2.0695 | 2.059 | 2.03 | 2.038 | 2.09 | 2.03
| Mean |  | 2.0567 | 2.0495 | 1.79 | 2.0517 | 2.0605 | 2.0685 | 1.8845 | 2.0602 | 2.0145 | 2.0185 | 2.1 | 2.0333 | 1.941 | 2.0317 | 2.05 | 2.0921 | 2.049 | 2.0267 | 2.0357 | 2.09 | 2.005
| Median |  | 2.065 | 2.0481 | 1.785 | 2.05 | 2.0594 | 2.063 | 1.901 | 2.063 | 2.009 | 2.0335 | 2.1 | 2 | 1.9608 | 2.035 | 2.05 | 2.1209 | 2.05 | 2.03 | 2.0355 | 2.09 | 1.99
| Std Dev. |  | 0.040332 | 0.016383 | 0.016733 | 0.01169 | 0.016295 | 0.019947 | 0.03453 | 0.032096 | 0.010075 | 0.042364 | 0 | 0.05164 | 0.036657 | 0.01472 | 0.054772 | 0.078484 | 0.013236 | 0.026583 | 0.0022509 | 0.014142 | 0.028107
| Rel.Std.Dev. |  | 1.961% | 0.7994% | 0.9348% | 0.5698% | 0.7908% | 0.9643% | 1.832% | 1.558% | 0.5001% | 2.099% | 0% | 2.54% | 1.889% | 0.7245% | 2.672% | 3.751% | 0.646% | 1.312% | 0.1106% | 0.6767% | 1.402%
| PDM3 |  | 0.491% | 0.1396% | -12.54% | 0.2467% | 0.6784% | 1.069% | -7.921% | 0.662% | -1.569% | -1.374% | 2.608% | -0.6491% | -5.16% | -0.7305% | 0.1653% | 2.222% | 0.1164% | -0.9748% | -0.5351% | 2.12% | -2.033%
| Z-Score (Absolute) |  | 0.26 | 0.01 | 9.02 | 0.09 | 0.4 | 0.67 | 5.73 | 0.39 | 1.2 | 1.06 | 1.77 | 0.55 | 3.76 | 0.61 | 0.03 | 1.5 | 0 | 0.78 | 0.47 | 1.42 | 1.53
| Analytical results for La in OREAS 313 (Certified Value 31.8 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | [9 empty] | ppm
|  |  | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 29.85 | 32.6 | 32.4 | 30.264 | 31.1 | 31.1 | 30 | 32.5 | 32 | 26 | 32 | 31.96 | [9 empty] | 1
| 1 | 2 | 28.85 | 33 | 33.2 | 27.9222 | 33.9 | 32.3 | 32 | 33.6 | 32 | 25 | 32.3 | 32.2 | [9 empty] | 12
| 1 | 3 | 28.93 | 33.9 | 32.6 | 30.5184 | 33.7 | 30.9 | 33 | 33.2 | 31 | 26 | 33.4 | 32.32 | [9 empty] | 16
| 1 | 4 | 30.13 | 33 | 33.1 | 31.5991 | 34.2 | 31.7 | 33 | 32.4 | 32 | 25 | 32.3 | 31.67 | [9 empty] | 31.7944
| 1 | 5 | 29.32 | 33.8 | 33.7 | 27.0931 | 32.6 | 31.3 | 31 | 33.1 | 33 | 25 | 31.9 | 32.97 | [9 empty] | 161
| 1 | 6 | 29.9 | 32.5 | 33.3 | 27.8942 | 30.2 | 31.7 | 29 | 33.2 | 32 | 25 | 32.5 | 32.04
| Mean |  | 29.4967 | 33.1333 | 33.05 | 29.2152 | 32.6167 | 31.5 | 31.3333 | 33 | 32 | 25.3333 | 32.4 | 32.1933
| Median |  | 29.585 | 33 | 33.15 | 29.0931 | 33.15 | 31.5 | 31.5 | 33.15 | 32 | 25 | 32.3 | 32.12
| Std Dev. |  | 0.54021 | 0.59217 | 0.47645 | 1.8112 | 1.6412 | 0.50596 | 1.633 | 0.46043 | 0.63246 | 0.5164 | 0.53666 | 0.44071
| Rel.Std.Dev. |  | 1.831% | 1.787% | 1.442% | 6.199% | 5.032% | 1.606% | 5.212% | 1.395% | 1.976% | 2.038% | 1.656% | 1.369%
| PDM3 |  | -7.227% | 4.211% | 3.949% | -8.112% | 2.586% | -0.926% | -1.45% | 3.792% | 0.6466% | -20.32% | 1.905% | 1.255%
| Z-Score (Absolute) |  | 2.1 | 0.84 | 0.77 | 2.33 | 0.42 | 0.48 | 0.62 | 0.73 | 0.08 | 5.47 | 0.25 | 0.08
| Analytical results for Li in OREAS 313 (Certified Value 23.6 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.02 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [10 empty] | 1
2
| 1 | 1 | 23 | 41.8035 | 21 | 23 | 27.4487 | 29 | 23 | 26 | 22 | 24 | 22 | [10 empty] | 1
| 1 | 2 | 23 | 40.1375 | 24 | 23 | 28.5933 | 30 | 35 | 26 | 22 | 24 | 20 | [10 empty] | #N/A
| 1 | 3 | 23 | 42.1095 | 23 | 23 | 23.3074 | 50 | 23 | 25 | 22 | 25 | 23 | [10 empty] | 16
| 1 | 4 | 24 | 41.3615 | 23 | 23 | 22.0386 | 54 | 23 | 27 | 22 | 24 | 23 | [10 empty] | 23.5877
| 1 | 5 | 23 | 41.4125 | 23 | 23 | 25.1158 | 26 | 24 | 25 | 22 | 23 | 23 | [10 empty] | 162
| 1 | 6 | 23 | 42.6195 | 22 | 22 | 27.9329 | 29 | 31 | 25 | 23 | 24 | 23
| Mean |  | 23.1667 | 41.574 | 22.6667 | 22.8333 | 25.7395 | 36.3333 | 26.5 | 25.6667 | 22.1667 | 24 | 22.3333
| Median |  | 23 | 41.608 | 23 | 23 | 26.2822 | 29.5 | 23.5 | 25.5 | 22 | 24 | 23
| Std Dev. |  | 0.40825 | 0.84466 | 1.0328 | 0.40825 | 2.6787 | 12.2746 | 5.2058 | 0.8165 | 0.40825 | 0.63246 | 1.2111
| Rel.Std.Dev. |  | 1.762% | 2.032% | 4.556% | 1.788% | 10.41% | 33.78% | 19.64% | 3.181% | 1.842% | 2.635% | 5.423%
| PDM3 |  | -1.785% | 76.25% | -3.905% | -3.198% | 9.122% | 54.03% | 12.35% | 8.814% | -6.025% | 1.748% | -5.318%
| Z-Score (Absolute) |  | 0.34 | 7.11 | 0.54 | 0.47 | 0.7 | 4.99 | 1.01 | 0.67 | 0.74 | 0 | 0.67
| Analytical results for Lu in OREAS 313 (Certified Value 0.3 ppm) | [22 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [13 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [13 empty] | 2
3
| 1 | 1 | < 0.5 | 0.3 | 0.23623 | 0.3 | < 2 | 0.3 | 0.34 | 0.32 | [13 empty] | 1
| 1 | 2 | < 0.5 | 0.3 | 0.25674 | 0.29 | < 2 | 0.3 | 0.28 | 0.31 | [13 empty] | 5
| 1 | 3 | < 0.5 | 0.3 | 0.24337 | 0.28 | < 2 | 0.3 | 0.32 | 0.32 | [13 empty] | 16
| 1 | 4 | < 0.5 | 0.3 | 0.24666 | 0.3 | < 2 | 0.3 | 0.34 | 0.32 | [13 empty] | 0.30467
| 1 | 5 | < 0.5 | 0.3 | 0.24326 | 0.3 | < 2 | 0.3 | 0.28 | 0.32 | [13 empty] | 163
| 1 | 6 | < 0.5 | 0.3 | 0.24922 | 0.29 | < 2 | 0.3 | 0.32 | 0.31
| Mean |  |  | 0.3 | 0.24591 | 0.29333 |  | 0.3 | 0.31333 | 0.31667
| Median |  |  | 0.3 | 0.24501 | 0.295 |  | 0.3 | 0.32 | 0.32
| Std Dev. |  |  | 0 | 0.0068686 | 0.008165 |  | 0 | 0.027325 | 0.005164
| Rel.Std.Dev. |  |  | 0% | 2.793% | 2.784% |  | 0% | 8.721% | 1.631%
| PDM3 |  |  | -1.532% | -19.28% | -3.72% |  | -1.532% | 2.845% | 3.939%
| Z-Score (Absolute) |  | 2.25 | 0 | 2.43 | 0.3 | 31.47 | 0 | 0.6 | 0.75
| Analytical results for Mg in OREAS 313 (Certified Value 1.05 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 1 | 1.0579 | 1.03 | 1.06 | 1.0563 | 1.025 | 0.989 | 1.061 | 1.079 | 1.013 | 1.04 | 1.04 | 1.0431 | 1.08 | 1.09 | 1.098 | 1.02 | 1.027 | 1.07 | 1 |  | 1
| 1 | 2 | 1.01 | 1.0591 | 1.02 | 1.06 | 1.0538 | 1.013 | 1.049 | 1.037 | 1.037 | 1.049 | 1.06 | 1.02 | 1.0534 | 1.08 | 1.11 | 1.086 | 1.03 | 1.034 | 1.06 | 0.98 |  | #N/A
| 1 | 3 | 1 | 1.0611 | 1.05 | 1.06 | 1.061 | 1.031 | 1.061 | 1.019 | 1.061 | 1.049 | 1.07 | 1.03 | 1.0724 | 1.08 | 1.08 | 1.073 | 1.01 | 1.033 | 1.06 | 1 |  | 16
| 1 | 4 | 1.07 | 1.0714 | 1.04 | 1.06 | 1.0563 | 1.025 | 1.019 | 1.037 | 1.061 | 1.013 | 1.06 | 1.02 | 1.0686 | 1.09 | 1.1 | 1.086 | 1.01 | 1.028 | 1.05 | 1 |  | 1.0465
| 1 | 5 | 1 | 1.0713 | 1.02 | 1.06 | 1.0551 | 1.019 | 1.019 | 1.019 | 1.061 | 1.067 | 1.07 | 1.04 | 1.0452 | 1.07 | 1.07 | 1.079 | 1.01 | 1.034 | 1.06 | 1.04 |  | 164
| 1 | 6 | 1 | 1.0946 | 1.04 | 1.08 | 1.076 | 1.025 | 1.037 | 1.037 | 1.061 | 1.031 | 1.06 | 1.04 | 1.0628 | 1.08 | 1.09 | 1.086 | 1.01 | 1.028 | 1.06 | 1.03
| Mean |  | 1.0133 | 1.0692 | 1.0333 | 1.0633 | 1.0597 | 1.023 | 1.029 | 1.035 | 1.06 | 1.037 | 1.06 | 1.0317 | 1.0576 | 1.08 | 1.09 | 1.0847 | 1.015 | 1.0307 | 1.06 | 1.0083
| Median |  | 1 | 1.0662 | 1.035 | 1.06 | 1.0563 | 1.025 | 1.028 | 1.037 | 1.061 | 1.04 | 1.06 | 1.035 | 1.0581 | 1.08 | 1.09 | 1.086 | 1.01 | 1.0305 | 1.06 | 1
| Std Dev. |  | 0.028048 | 0.013776 | 0.012111 | 0.008165 | 0.0083262 | 0.0061968 | 0.025644 | 0.015492 | 0.013372 | 0.021799 | 0.010954 | 0.0098319 | 0.012225 | 0.0063246 | 0.014142 | 0.0083825 | 0.0083666 | 0.0033267 | 0.0063246 | 0.022286
| Rel.Std.Dev. |  | 2.768% | 1.288% | 1.172% | 0.7679% | 0.7857% | 0.6057% | 2.492% | 1.497% | 1.261% | 2.102% | 1.033% | 0.953% | 1.156% | 0.5856% | 1.297% | 0.7728% | 0.8243% | 0.3228% | 0.5967% | 2.21%
| PDM3 |  | -3.173% | 2.169% | -1.262% | 1.605% | 1.26% | -2.249% | -1.676% | -1.102% | 1.286% | -0.9112% | 1.286% | -1.421% | 1.056% | 3.198% | 4.153% | 3.643% | -3.013% | -1.516% | 1.286% | -3.65%
| Z-Score (Absolute) |  | 1.41 | 0.9 | 0.58 | 0.66 | 0.51 | 1 | 0.75 | 0.5 | 0.53 | 0.42 | 0.52 | 0.65 | 0.42 | 1.34 | 1.76 | 1.53 | 1.34 | 0.69 | 0.52 | 1.61
| Analytical results for Mn in OREAS 313 (Certified Value 0.188 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.183 | 0.1924 | 0.1787 | 0.191 | 0.18739 | 0.186 | 0.178 | 0.194 | 0.178 | 0.186 | 0.187 | 0.1869 | 0.20576 | 0.1927 | 0.2 | 0.194 | 0.1856 | 0.1757 | 0.1962 | 0.1801 |  | 1
| 1 | 2 | 0.183 | 0.19325 | 0.1801 | 0.191 | 0.18567 | 0.186 | 0.186 | 0.194 | 0.178 | 0.194 | 0.19 | 0.186 | 0.20506 | 0.1928 | 0.2 | 0.194 | 0.1853 | 0.1765 | 0.1952 | 0.1783 |  | 35
| 1 | 3 | 0.183 | 0.19277 | 0.1836 | 0.192 | 0.18812 | 0.186 | 0.186 | 0.186 | 0.178 | 0.194 | 0.191 | 0.1871 | 0.1942 | 0.1922 | 0.19 | 0.194 | 0.1847 | 0.1765 | 0.1959 | 0.1801 |  | 16
| 1 | 4 | 0.181 | 0.19245 | 0.1834 | 0.191 | 0.18592 | 0.194 | 0.178 | 0.186 | 0.178 | 0.186 | 0.185 | 0.1852 | 0.20321 | 0.1926 | 0.2 | 0.194 | 0.1855 | 0.1766 | 0.1952 | 0.1794 |  | 0.18803
| 1 | 5 | 0.183 | 0.19349 | 0.1803 | 0.191 | 0.18582 | 0.186 | 0.178 | 0.186 | 0.178 | 0.194 | 0.192 | 0.1884 | 0.2052 | 0.1907 | 0.2 | 0.194 | 0.1808 | 0.1752 | 0.1958 | 0.1886 |  | 165
| 1 | 6 | 0.182 | 0.19411 | 0.1801 | 0.194 | 0.18986 | 0.186 | 0.186 | 0.194 | 0.178 | 0.186 | 0.181 | 0.1883 | 0.19319 | 0.1934 | 0.2 | 0.194 | 0.1842 | 0.1751 | 0.1971 | 0.1861
| Mean |  | 0.1825 | 0.19308 | 0.18103 | 0.19167 | 0.18713 | 0.18733 | 0.182 | 0.19 | 0.178 | 0.19 | 0.18767 | 0.18698 | 0.2011 | 0.1924 | 0.19833 | 0.194 | 0.18435 | 0.17593 | 0.1959 | 0.1821
| Median |  | 0.183 | 0.19301 | 0.1802 | 0.191 | 0.18666 | 0.186 | 0.182 | 0.19 | 0.178 | 0.19 | 0.1885 | 0.187 | 0.20413 | 0.19265 | 0.2 | 0.194 | 0.185 | 0.1761 | 0.19585 | 0.1801
| Std Dev. |  | 0.00083666 | 0.00066357 | 0.0019957 | 0.0012111 | 0.0016614 | 0.003266 | 0.0043818 | 0.0043818 | 3.0405e-17 | 0.0043818 | 0.0041793 | 0.0012576 | 0.0058124 | 0.00091869 | 0.0040825 | 3.0405e-17 | 0.0018185 | 0.00068896 | 0.00070993 | 0.0041948
| Rel.Std.Dev. |  | 0.4584% | 0.3437% | 1.102% | 0.6319% | 0.8878% | 1.743% | 2.408% | 2.306% | 1.708e-14% | 2.306% | 2.227% | 0.6726% | 2.89% | 0.4775% | 2.058% | 1.567e-14% | 0.9864% | 0.3916% | 0.3624% | 2.304%
| PDM3 |  | -2.94% | 2.687% | -3.72% | 1.935% | -0.4778% | -0.3695% | -3.206% | 1.049% | -5.333% | 1.049% | -0.1922% | -0.5556% | 6.954% | 2.325% | 5.481% | 3.176% | -1.956% | -6.432% | 4.187% | -3.153%
| Z-Score (Absolute) |  | 0.64 | 0.74 | 0.84 | 0.56 | 0.04 | 0.03 | 0.71 | 0.3 | 1.22 | 0.3 | 0.03 | 0.06 | 1.79 | 0.65 | 1.43 | 0.81 | 0.4 | 1.5 | 1.11 | 0.7
| Analytical results for Mo in OREAS 313 (Certified Value 5.14 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.12 | 2.13 | 2.14 | 2.16 | 2.20 | 2.21 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | 5 | 4.6 | 6 | 6 | 5 | 4.8235 | 6 | 6 | < 10 | 7 | < 10 | [10 empty] | 1
| 1 | 2 | 5 | 4.7 | 6 | 9 | 5 | 4.9398 | 6 | 6 | < 10 | 8 | < 10 | [10 empty] | 14
| 1 | 3 | 5 | 4.7 | 5 | 7 | 5 | 4.7632 | 9 | 4 | < 10 | 8 | < 10 | [10 empty] | 16
| 1 | 4 | 5 | 4.8 | 5 | 7 | 5 | 4.5783 | 6 | 5 | < 10 | 8 | < 10 | [10 empty] | 5.1378
| 1 | 5 | 5 | 4.3 | 6 | 7 | 5 | 4.5601 | 6 | 5 | < 10 | 8 | < 10 | [10 empty] | 166
| 1 | 6 | 5 | 4.6 | 6 | 6 | 5 | 4.423 | 6 | 4 | < 10 | 7 | < 10
| Mean |  | 5 | 4.6167 | 5.6667 | 7 | 5 | 4.6813 | 6.5 | 5 |  | 7.6667
| Median |  | 5 | 4.65 | 6 | 7 | 5 | 4.6707 | 6 | 5 |  | 8
| Std Dev. |  | 0 | 0.17224 | 0.5164 | 1.0954 | 0 | 0.19274 | 1.2247 | 0.89443 |  | 0.5164
| Rel.Std.Dev. |  | 0% | 3.731% | 9.113% | 15.65% | 0% | 4.117% | 18.84% | 17.89% |  | 6.736%
| PDM3 |  | -2.682% | -10.14% | 10.29% | 36.24% | -2.682% | -8.885% | 26.51% | -2.682% |  | 49.22%
| Z-Score (Absolute) |  | 0 | 0.81 | 1.41 | 4.23 | 0 | 0.67 | 3.17 | 0 | 0 | 5.64 | 0
| Analytical results for Na in OREAS 313 (Indicative Value 0.093 wt.%) | [22 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.04 | 2.05 | 2.19 | [18 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES/MS | [18 empty] | 3
3
| 1 | 1 | 0.12 | 0.08 | 0.074 | [18 empty] | 1
| 1 | 2 | 0.1 | 0.08 | 0.074 | [18 empty] | 6
| 1 | 3 | 0.13 | 0.08 | 0.067 | [18 empty] | 16
| 1 | 4 | 0.16 | 0.08 | 0.074 | [18 empty] | 0.093348
| 1 | 5 | 0.15 | 0.08 | 0.067 | [18 empty] | 32
| 1 | 6 | 0.11 | 0.08 | 0.074
| Mean |  | 0.12833 | 0.08 | 0.071667
| Median |  | 0.125 | 0.08 | 0.074
| Std Dev. |  | 0.023166 | 0 | 0.0036148
| Rel.Std.Dev. |  | 18.05% | 0% | 5.044%
| PDM3 |  | 37.48% | -14.3% | -23.23%
| Z-Score (Absolute) |  | 3.93 | 0 | 0.67
| Analytical results for Nb in OREAS 313 (Certified Value 7.97 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.19 | 2.20 | 2.24 | [11 empty] | ppm
|  |  | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*OES/MS | [11 empty] | 2
3
| 1 | 1 | 185 | 7.4 | 8.8 | 7.8493 | 8.6 | 6 | 8.0251 | 7.6 | 20 | 8 | [11 empty] | 1
| 1 | 2 | 177.83 | 8.5 | 8.7 | 7.7968 | 8.2 | 7 | 7.0702 | 7.4 | 18 | 9 | [11 empty] | 16
| 1 | 3 | 186.08 | 8.5 | 8.6 | 7.6036 | 9.2 | 7 | 6.58 | 7.5 | 18 | 9 | [11 empty] | 16
| 1 | 4 | 187.94 | 8.2 | 8.5 | 7.6929 | 8.4 | 8 | 6.9729 | 7.5 | 21 | 10 | [11 empty] | 7.9674
| 1 | 5 | 179.74 | 6.3 | 8.6 | 7.7215 | 8.6 | 7 | 7.8972 | 7.4 | 18 | 9 | [11 empty] | 167
| 1 | 6 | 193.3 | 8 | 8.8 | 7.7535 | 8.3 | 6 | 6.8451 | 7.6 | 20 | 8
| Mean |  | 184.9817 | 7.8167 | 8.6667 | 7.7363 | 8.55 | 6.8333 | 7.2317 | 7.5 | 19.1667 | 8.8333
| Median |  | 185.54 | 8.1 | 8.65 | 7.7375 | 8.5 | 7 | 7.0215 | 7.5 | 19 | 9
| Std Dev. |  | 5.6174 | 0.84715 | 0.12111 | 0.085355 | 0.35637 | 0.75277 | 0.58985 | 0.089443 | 1.3292 | 0.75277
| Rel.Std.Dev. |  | 3.037% | 10.84% | 1.397% | 1.103% | 4.168% | 11.02% | 8.156% | 1.193% | 6.935% | 8.522%
| PDM3 |  | 2222% | -1.892% | 8.776% | -2.902% | 7.312% | -14.23% | -9.234% | -5.867% | 140.6% | 10.87%
| Z-Score (Absolute) |  | 204.24 | 0 | 0.98 | 0.09 | 0.85 | NA | 0.67 | 0.37 | NA | NA
Results from laboratories 2.13, 2.20 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 313 (Certified Value 27.4 ppm) | [22 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 1
2
| 1 | 1 | 28.31 | 27.1 | 27.7166 | 27.6 | 27.3 | 23 | 26.7 | 25.5 | 26.5 | [12 empty] | 1
| 1 | 2 | 28.48 | 28.9 | 29.9729 | 31.3 | 28.1 | 25 | 27.2 | 25.3 | 26.7 | [12 empty] | 7
| 1 | 3 | 29.27 | 27.7 | 27.8088 | 31.7 | 27 | 26 | 26.7 | 26.3 | 27.1 | [12 empty] | 16
| 1 | 4 | 27.92 | 28.7 | 28.9999 | 27.6 | 27.4 | 27 | 26.2 | 26.1 | 26.9 | [12 empty] | 27.4076
| 1 | 5 | 28.97 | 28.5 | 27.8165 | 31.9 | 27.4 | 25 | 26.9 | 25.8 | 27.9 | [12 empty] | 168
| 1 | 6 | 28.18 | 28.3 | 29.4662 | 26.1 | 28 | 23 | 27 | 25.7 | 27
| Mean |  | 28.5217 | 28.2 | 28.6301 | 29.3667 | 27.5333 | 24.8333 | 26.7833 | 25.7833 | 27.0167
| Median |  | 28.395 | 28.4 | 28.4082 | 29.45 | 27.4 | 25 | 26.8 | 25.75 | 26.95
| Std Dev. |  | 0.5072 | 0.67823 | 0.98081 | 2.55 | 0.4274 | 1.6021 | 0.34303 | 0.37103 | 0.48339
| Rel.Std.Dev. |  | 1.778% | 2.405% | 3.426% | 8.683% | 1.552% | 6.451% | 1.281% | 1.439% | 1.789%
| PDM3 |  | 4.065% | 2.891% | 4.461% | 7.148% | 0.4587% | -9.393% | -2.278% | -5.926% | -1.426%
| Z-Score (Absolute) |  | 0.67 | 0.45 | 0.75 | 1.25 | 0 | 1.84 | 0.51 | 1.19 | 0.35
| Analytical results for Ni in OREAS 313 (Certified Value 19.3 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 1
1
| 1 | 1 | < 20 | 16.67 | < 20 | 30 | 190 | 20 | 20 | 40 | 19 | < 20 |  | < 50 | < 20 | 16 | 15 | 62 | < 50 |  |  |  |  | 1
| 1 | 2 | < 20 | 15.25 | < 20 | 30 | 60 | 20 | 20 | 30 | 19 | < 20 | 22 | < 50 | < 20 | 20 | 14 | 20 | < 50 |  |  |  |  | 18
| 1 | 3 | < 20 | 15.87 | < 20 | 30 | 50 | 20 | 30 | 20 | 21 | < 20 | 26 | < 50 | < 20 | 19 | 15 | 40 | < 50 |  |  |  |  | 16
| 1 | 4 | < 20 | 15.57 | 20 | 40 | 50 | < 20 | 20 | 30 | 20 | < 20 |  | < 50 | 22 | 19 | 15 | 47 | < 50 |  |  |  |  | 19.305
| 1 | 5 | < 20 | 16.45 | < 20 | 30 | 70 | < 20 | 20 | 30 | 21 | < 20 |  | < 50 | < 20 | 19 | 15 | 126 | < 50 |  |  |  |  | 169
| 1 | 6 | < 20 | 15.17 | < 20 | 30 | 40 | 50 | 20 | 30 | 21 | < 20 | 26 | < 50 | < 20 | 17 | 15 | 52 | < 50
| Mean |  |  | 15.83 | 20 | 31.6667 | 76.6667 | 27.5 | 21.6667 | 30 | 20.1667 |  | 24.6667 |  | 22 | 18.3333 | 14.8333 | 57.8333
| Median |  |  | 15.72 | 20 | 30 | 55 | 20 | 20 | 30 | 20.5 |  | 26 |  | 22 | 19 | 15 | 49.5
| Std Dev. |  |  | 0.62148 |  | 4.0825 | 56.4506 | 15 | 4.0825 | 6.3246 | 0.98319 |  | 2.3094 |  |  | 1.5055 | 0.40825 | 36.2349
| Rel.Std.Dev. |  |  | 3.926% |  | 12.89% | 73.63% | 54.55% | 18.84% | 21.08% | 4.875% |  | 9.362% |  |  | 8.212% | 2.752% | 62.65%
| PDM3 |  |  | -18% | 3.6% | 64.03% | 297.1% | 42.45% | 12.23% | 55.4% | 4.463% |  | 27.77% |  | 13.96% | -5.033% | -23.16% | 199.6%
| Z-Score (Absolute) |  | 0.89 | 0.27 | NA | NA | NA | NA | NA | NA | 0.2 | 0.89 | 0.67 | 0.71 | 0.67 | 0 | 0.37 | 4.21 | 0.71
Results from laboratories 2.07, 2.08, 2.09, 2.10, 2.11 and 2.12 were removed due to their 10 ppm reading resolution.
| Analytical results for P in OREAS 313 (Certified Value 0.031 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.13 | 2.19 | 2.20 | 2.21 | [12 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.03 | 0.031866 | 0.03 | 0.03 | 0.035992 | 0.03 | 0.035 | 0.027 | 0.032 | [12 empty] | 1
| 1 | 2 | 0.03 | 0.031953 | 0.03 | 0.03 | 0.037387 | 0.03 | 0.031 | 0.027 | 0.038 | [12 empty] | #N/A
| 1 | 3 | 0.03 | 0.032731 | 0.03 | 0.03 | 0.035501 | 0.03 | 0.035 | 0.0272 | 0.031 | [12 empty] | 16
| 1 | 4 | 0.03 | 0.031476 | 0.03 | 0.03 | 0.035302 | 0.03 | 0.031 | 0.0251 | 0.038 | [12 empty] | 0.031288
| 1 | 5 | 0.03 | 0.032792 | 0.03 | 0.03 | 0.035261 | 0.03 | 0.031 | 0.0287 | 0.037 | [12 empty] | 170
| 1 | 6 | 0.03 | 0.031225 | 0.03 | 0.03 | 0.036479 | 0.03 | 0.031 | 0.0266 | 0.032
| Mean |  | 0.03 | 0.032007 | 0.03 | 0.03 | 0.035987 | 0.03 | 0.032333 | 0.026933 | 0.034667
| Median |  | 0.03 | 0.03191 | 0.03 | 0.03 | 0.035746 | 0.03 | 0.031 | 0.027 | 0.0345
| Std Dev. |  | 0 | 0.00064136 | 0 | 0 | 0.00082881 | 0 | 0.0020656 | 0.0011553 | 0.0033267
| Rel.Std.Dev. |  | 0% | 2.004% | 0% | 0% | 2.303% | 0% | 6.388% | 4.289% | 9.596%
| PDM3 |  | -4.118% | 2.297% | -4.118% | -4.118% | 15.02% | -4.118% | 3.34% | -13.92% | 10.8%
| Z-Score (Absolute) |  | 0 | 0.68 | 0 | 0 | 2.02 | 0 | 0.67 | 1.03 | 1.57
| Analytical results for Pb in OREAS 313 (Certified Value 3.77 wt.%) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 3.71 | 3.7366 | > 1 | 3.75 | 3.59 | 3.77 | 3.81 | 3.56 | > 0.5 | 3.87 | 3.7892 | 3.9576 | 3.9 | 3.85 | 3.7677 | 3.7807 | > 2.5 |  |  |  |  | 1
| 1 | 2 | 3.63 | 3.7378 | > 1 | 3.66 | 3.76 | 3.74 | 3.78 | 3.69 | > 0.5 | 3.84 | 3.7948 | 4.0966 | 4.04 | 3.91 | 3.7014 | 3.7749 | > 2.5 |  |  |  |  | 40
| 1 | 3 | 3.77 | 3.7289 | > 1 | 3.74 | 3.8 | 3.71 | 3.78 | 3.71 | > 0.5 | 3.86 | 3.8071 | 4.047 | 3.83 | 3.83 | 3.721 | 3.7802 | > 2.5 |  |  |  |  | 16
| 1 | 4 | 3.73 | 3.7132 | > 1 | 3.71 | 3.7 | 3.8 | 3.78 | 3.58 | > 0.5 | 3.9 | 3.8187 | 4.0631 | 3.89 | 3.91 | 3.7663 | 3.7897 | > 2.5 |  |  |  |  | 3.7712
| 1 | 5 | 3.75 | 3.7496 | > 1 | 3.74 | 3.62 | 3.72 | 3.81 | 3.77 | > 0.5 | 3.88 | 3.8132 | 3.9953 | 3.82 | 3.87 | 3.7595 | 3.7958 | > 2.5 |  |  |  |  | 171
| 1 | 6 | 3.73 | 3.7312 | > 1 | 3.73 | 3.74 | 3.75 | 3.81 | 3.71 | > 0.5 | 3.85 | 3.779 | 3.9995 | 3.95 | 3.86 | 3.7351 | 3.743 | > 2.5
| Mean |  | 3.72 | 3.7329 |  | 3.7217 | 3.7017 | 3.7483 | 3.795 | 3.67 |  | 3.8667 | 3.8003 | 4.0265 | 3.905 | 3.8717 | 3.7418 | 3.7774
| Median |  | 3.73 | 3.7339 |  | 3.735 | 3.72 | 3.745 | 3.795 | 3.7 |  | 3.865 | 3.8009 | 4.0232 | 3.895 | 3.865 | 3.7473 | 3.7805
| Std Dev. |  | 0.04858 | 0.01203 |  | 0.033116 | 0.082077 | 0.033116 | 0.016432 | 0.082219 |  | 0.021602 | 0.015204 | 0.051232 | 0.08167 | 0.032506 | 0.027181 | 0.018429
| Rel.Std.Dev. |  | 1.306% | 0.3223% |  | 0.8898% | 2.217% | 0.8835% | 0.433% | 2.24% |  | 0.5587% | 0.4001% | 1.272% | 2.091% | 0.8396% | 0.7264% | 0.4879%
| PDM3 |  | -1.357% | -1.016% |  | -1.313% | -1.843% | -0.606% | 0.6315% | -2.683% |  | 2.532% | 0.7726% | 6.771% | 3.548% | 2.664% | -0.7783% | 0.1643%
| Z-Score (Absolute) |  | 0.69 | 0.48 | NA | 0.66 | 0.98 | 0.23 | 0.52 | 1.49 | NA | 1.67 | 0.6 | 4.23 | 2.28 | 1.75 | 0.34 | 0.23 | NA
| Analytical results for Pr in OREAS 313 (Certified Value 7.46 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 7.5 | 7.3 | 6.9503 | 7.1 | 7.48 | 6 | 7.19 | 7.17 | 7.04 | [12 empty] | 1
| 1 | 2 | 7.6 | 7.9 | 7.209 | 8.7 | 7.67 | 7 | 7.29 | 7.41 | 7.27 | [12 empty] | 19
| 1 | 3 | 7.8 | 7.7 | 6.9985 | 8.5 | 7.27 | 7 | 7.26 | 7.61 | 7.16 | [12 empty] | 16
| 1 | 4 | 7.5 | 7.9 | 7.1649 | 7.8 | 7.45 | 7 | 7.31 | 7.3 | 6.96 | [12 empty] | 7.4635
| 1 | 5 | 7.7 | 7.7 | 7.1316 | 8.1 | 7.43 | 7 | 7.32 | 7.22 | 7.32 | [12 empty] | 172
| 1 | 6 | 7.5 | 8 | 7.315 | 8.3 | 7.47 | 6 | 7.42 | 7.41 | 7.19
| Mean |  | 7.6 | 7.75 | 7.1282 | 8.0833 | 7.4617 | 6.6667 | 7.2983 | 7.3533 | 7.1567
| Median |  | 7.55 | 7.8 | 7.1482 | 8.2 | 7.46 | 7 | 7.3 | 7.355 | 7.175
| Std Dev. |  | 0.12649 | 0.251 | 0.13508 | 0.57417 | 0.12781 | 0.5164 | 0.075741 | 0.15908 | 0.13633
| Rel.Std.Dev. |  | 1.664% | 3.239% | 1.895% | 7.103% | 1.713% | 7.746% | 1.038% | 2.163% | 1.905%
| PDM3 |  | 1.829% | 3.838% | -4.493% | 8.304% | -0.02492% | -10.68% | -2.213% | -1.476% | -4.111%
| Z-Score (Absolute) |  | 0.59 | 1.04 | 0.85 | 2.06 | 0.16 | NA | 0.33 | 0.16 | 0.76
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 313 (Certified Value 109 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.22 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 102 | 113.1 | 105.7473 | 111 | 111 | 117.9832 | 111 | 107 | 114.3 | 103 | 107 | 102 | [9 empty] | 1
| 1 | 2 | 107 | 115 | 107.6977 | 110 | 111 | 115.7403 | 112.6 | 143 | 111.3 | 105 | 112.1 | 100 | [9 empty] | 20
| 1 | 3 | 105 | 112.3 | 105.3646 | 109 | 111 | 120.9207 | 114.2 | 153 | 112.8 | 107 | 109.8 | 102 | [9 empty] | 16
| 1 | 4 | 104 | 112.2 | 109.2269 | 109 | 110 | 116.9201 | 112 | 108 | 113.3 | 104 | 110.5 | 96 | [9 empty] | 109.3781
| 1 | 5 | 103 | 110.8 | 106.6142 | 109 | 111 | 120.1847 | 111.8 | 243 | 111.3 | 106 | 111.2 | 106 | [9 empty] | 173
| 1 | 6 | 105 | 110.9 | 106.048 | 107 | 113 | 121.4055 | 110.4 | 191 | 111.9 | 104 | 109.3 | 101
| Mean |  | 104.3333 | 112.3833 | 106.7831 | 109.1667 | 111.1667 | 118.8591 | 112 | 157.5 | 112.4833 | 104.8333 | 109.9833 | 101.1667
| Median |  | 104.5 | 112.25 | 106.3311 | 109 | 111 | 119.084 | 111.9 | 148 | 112.35 | 104.5 | 110.15 | 101.5
| Std Dev. |  | 1.7512 | 1.5562 | 1.4466 | 1.3292 | 0.98319 | 2.3128 | 1.3266 | 52.2752 | 1.2007 | 1.472 | 1.768 | 3.2506
| Rel.Std.Dev. |  | 1.678% | 1.385% | 1.355% | 1.218% | 0.8844% | 1.946% | 1.185% | 33.19% | 1.067% | 1.404% | 1.607% | 3.213%
| PDM3 |  | -4.612% | 2.748% | -2.372% | -0.1933% | 1.635% | 8.668% | 2.397% | 44% | 2.839% | -4.155% | 0.5534% | -7.507%
| Z-Score (Absolute) |  | 1.48 | 0.43 | 0.9 | 0.33 | 0.14 | 1.96 | 0.34 | 11.1 | 0.45 | 1.36 | 0.14 | 2.23
| Analytical results for Re in OREAS 313 (Indicative Value < 0.1 ppm) | [22 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.01 | 2.14 | 2.24 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | [18 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 0.01 | [18 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.01 | [18 empty] | 27
| 1 | 3 | < 0.1 | < 0.1 | < 0.01 | [18 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.01 | [18 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 0.01 | [18 empty] | 33
| 1 | 6 | < 0.1 | < 0.1 | < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 313 (Certified Value 4.57 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.16 | 2.20 | 2.23 | 2.24 | [8 empty] | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 4.56 | 6.77 | 4.67 | 4.44 | 4.66 | 4.51 | 4.63 | 4.5 | 4.4975 | 4.53 | 4.64 | 4.56 | 4.42 | [8 empty] | 1
| 1 | 2 | 4.59 | 6.98 | 4.64 | 4.59 | 4.67 | 4.53 | 4.62 | 4.42 | 4.666 | 4.52 | 4.56 | 4.55 | 4.35 | [8 empty] | #N/A
| 1 | 3 | 4.54 | 6.92 | 4.69 | 4.68 | 4.57 | 4.53 | 4.72 | 4.55 | 4.6396 | 4.58 | 4.64 | 4.58 | 4.38 | [8 empty] | 16
| 1 | 4 | 4.52 | 6.71 | 4.66 | 4.54 | 4.66 | 4.54 | 4.65 | 4.48 | 4.6298 | 4.59 | 4.67 | 4.56 | 4.42 | [8 empty] | 4.5735
| 1 | 5 | 4.54 | 6.72 | 4.67 | 4.5 | 4.65 | 4.5 | 4.72 | 4.5 | 4.5247 | 4.51 | 4.58 | 4.52 | 4.64 | [8 empty] | 174
| 1 | 6 | 4.55 | 6.98 | 4.68 | 4.6 | 4.56 | 4.53 | 4.7 | 4.44 | 4.6644 | 4.55 | 4.67 | 4.55 | 4.6
| Mean |  | 4.55 | 6.8467 | 4.6683 | 4.5583 | 4.6283 | 4.5233 | 4.6733 | 4.4817 | 4.6037 | 4.5467 | 4.6267 | 4.5533 | 4.4683
| Median |  | 4.545 | 6.845 | 4.67 | 4.565 | 4.655 | 4.53 | 4.675 | 4.49 | 4.6347 | 4.54 | 4.64 | 4.555 | 4.42
| Std Dev. |  | 0.023664 | 0.1277 | 0.017224 | 0.084004 | 0.049565 | 0.015055 | 0.045461 | 0.046655 | 0.07356 | 0.03266 | 0.046332 | 0.019664 | 0.12106
| Rel.Std.Dev. |  | 0.5201% | 1.865% | 0.369% | 1.843% | 1.071% | 0.3328% | 0.9728% | 1.041% | 1.598% | 0.7183% | 1.001% | 0.4319% | 2.709%
| PDM3 |  | -0.5138% | 49.7% | 2.074% | -0.3316% | 1.199% | -1.097% | 2.183% | -2.008% | 0.6597% | -0.5867% | 1.162% | -0.441% | -2.299%
| Z-Score (Absolute) |  | 0.08 | 22.58 | 1.09 | 0 | 0.69 | 0.35 | 1.13 | 0.76 | 0.45 | 0.12 | 0.67 | 0.05 | 0.89
| Analytical results for Sb in OREAS 313 (Certified Value 48.5 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.12 | 2.13 | 2.14 | 2.16 | 2.21 | 2.23 | 2.24 | [12 empty] | ppm
|  |  | PF*OES | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [12 empty] | 1
2
| 1 | 1 | 46.6 | 53.3 | 48 | 41.1 | 32.1552 | 49.3 | < 50 | 47.6 | 49.76 | [12 empty] | 1
| 1 | 2 | 45.2 | 53 | 52 | 42.5 | 36.9465 | 47.7 | < 50 | 50 | 51.8 | [12 empty] | 21
| 1 | 3 | 44.9 | 51.6 | 49 | 43.9 | 37.5927 | 47.7 | < 50 | 48.9 | 51.99 | [12 empty] | 16
| 1 | 4 | 47.3 | 51.9 | 48 | 46.7 | 39.2821 | 50.9 | < 50 | 49 | 49.24 | [12 empty] | 48.5471
| 1 | 5 | 44.9 | 53.7 | 51 | 42.1 | 39.8666 | 49.3 | < 50 | 48.9 | 53.85 | [12 empty] | 175
| 1 | 6 | 46.8 | 52.6 | 47 | 40.9 | 31.5614 | 49.2 | < 50 | 48.3 | 51.54
| Mean |  | 45.95 | 52.6833 | 49.1667 | 42.8667 | 36.2341 | 49.0167 |  | 48.7833 | 51.3633
| Median |  | 45.9 | 52.8 | 48.5 | 42.3 | 37.2696 | 49.25 |  | 48.9 | 51.67
| Std Dev. |  | 1.071 | 0.81343 | 1.9408 | 2.1667 | 3.5584 | 1.2007 |  | 0.79854 | 1.666
| Rel.Std.Dev. |  | 2.331% | 1.544% | 3.947% | 5.055% | 9.82% | 2.45% |  | 1.637% | 3.244%
| PDM3 |  | -5.35% | 8.52% | 1.276% | -11.7% | -25.36% | 0.9672% |  | 0.4865% | 5.801%
| Z-Score (Absolute) |  | 0.67 | 0.93 | 0.09 | 1.41 | 2.99 | 0.06 | 5.66 | 0 | 0.61
| Analytical results for Sc in OREAS 313 (Certified Value 8.38 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.13 | 2.14 | 2.18 | 2.19 | 2.20 | 2.21 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*OES | BF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES | PF*OES | [12 empty] | 2
3
| 1 | 1 | < 20 | 7.6 | 8 | 8 | < 20 | 8 | 9 | 10 | < 5 | [12 empty] | 1
| 1 | 2 | < 20 | 7.4 | 8 | 8 | < 20 | 9 | 9 | 9 | < 5 | [12 empty] | 22
| 1 | 3 | < 20 | 7.5 | 8 | 8 | < 20 | 9 | 9 | 9 | < 5 | [12 empty] | 16
| 1 | 4 | < 20 | 7.4 | 8 | 8 | < 20 | 9 | 9 | 9 | < 5 | [12 empty] | 8.3806
| 1 | 5 | < 20 | 7.4 | 8 | 8 | < 20 | 9 | 9 | 9 | < 5 | [12 empty] | 176
| 1 | 6 | < 20 | 7.4 | 8 | 8 | < 20 | 8 | 9 | 9 | < 5
| Mean |  |  | 7.45 | 8 | 8 |  | 8.6667 | 9 | 9.1667
| Median |  |  | 7.4 | 8 | 8 |  | 9 | 9 | 9
| Std Dev. |  |  | 0.083666 | 0 | 0 |  | 0.5164 | 0 | 0.40825
| Rel.Std.Dev. |  |  | 1.123% | 0% | 0% |  | 5.958% | 0% | 4.454%
| PDM3 |  |  | -11.1% | -4.541% | -4.541% |  | 3.414% | 7.391% | 9.38%
| Z-Score (Absolute) |  | 1.35 | 1.23 | 0.67 | 0.67 | 1.35 | 0 | 0.34 | 0.51 | 6.24
| Analytical results for Se in OREAS 313 (Indicative Value < 10 ppm) | [22 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.12 | 2.14 | 2.24 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | [18 empty] | 1
1
| 1 | 1 | < 8 | < 20 | < 10 | [18 empty] | 1
| 1 | 2 | < 8 | < 20 | < 10 | [18 empty] | 2
| 1 | 3 | < 8 | < 20 | < 10 | [18 empty] | 16
| 1 | 4 | < 8 | < 20 | < 10 | [18 empty] | < 10
| 1 | 5 | < 8 | < 20 | < 10 | [18 empty] | 34
| 1 | 6 | < 8 | < 20 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Si in OREAS 313 (Certified Value 27.25 wt.%) | [22 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 25.4052 | 26.95 | 27.58 | 25.8 | 28.98 | 26.79 | 26.6 | 27.3 | 24.1 | 23.3029 | 27.3 | 27.8 | 26.851 | 27.82 | 27.31 | 28.4 | 25.8 |  |  |  |  | 1
| 1 | 2 | 25.5882 | 27.09 | 27.3 | 27.11 | 28.98 | 26.69 | 27.3 | 27.5 | 25.2 | 23.1691 | 27.2 | 28.5 | 26.814 | 28.15 | 27.454 | 28 | 25.7 |  |  |  |  | #N/A
| 1 | 3 | 25.6965 | 27.07 | 27.58 | 27.39 | 28.52 | 26.79 | 27.39 | 27.9 | 26.9 | 24.7585 | 27.3 | 27.5 | 26.846 | 27.66 | 27.199 | 28.2 | 26.2 |  |  |  |  | 16
| 1 | 4 | 25.5633 | 27.12 | 27.58 | 26.69 | 28.89 | 26.88 | 26.6 | 27.1 | 28 | 24.7156 | 27.3 | 28 | 26.912 | 27.78 | 27.978 | 28 | 25.8 |  |  |  |  | 27.2539
| 1 | 5 | 25.2033 | 27.01 | 27.58 | 26.6 | 28.61 | 26.79 | 27.91 | 27.1 | 26.4 | 22.4341 | 27.1 | 27.7 | 27.066 | 27.18 | 27.456 | 28 | 27.1 |  |  |  |  | 177
| 1 | 6 | 25.5892 | 27.29 | 27.58 | 27.3 | 28.89 | 26.79 | 27.11 | 25.5 | 23.1 | 23.8399 | 27.4 | 27.8 | 26.814 | 27.58 | 27.696 | 28.1 | 26.5
| Mean |  | 25.5076 | 27.0883 | 27.5333 | 26.815 | 28.8117 | 26.7883 | 27.1517 | 27.0667 | 25.6167 | 23.7034 | 27.2667 | 27.8833 | 26.8838 | 27.695 | 27.5155 | 28.1167 | 26.1833
| Median |  | 25.5757 | 27.08 | 27.58 | 26.9 | 28.89 | 26.79 | 27.205 | 27.2 | 25.8 | 23.5714 | 27.3 | 27.8 | 26.8485 | 27.72 | 27.455 | 28.05 | 26
| Std Dev. |  | 0.17608 | 0.116 | 0.11431 | 0.59075 | 0.19732 | 0.060139 | 0.50285 | 0.82381 | 1.8302 | 0.91798 | 0.10328 | 0.34303 | 0.096167 | 0.31923 | 0.2815 | 0.16021 | 0.54191
| Rel.Std.Dev. |  | 0.6903% | 0.4282% | 0.4152% | 2.203% | 0.6849% | 0.2245% | 1.852% | 3.044% | 7.145% | 3.873% | 0.3788% | 1.23% | 0.3577% | 1.153% | 1.023% | 0.5698% | 2.07%
| PDM3 |  | -6.408% | -0.6075% | 1.025% | -1.61% | 5.716% | -1.708% | -0.3752% | -0.687% | -6.007% | -13.03% | 0.0468% | 2.309% | -1.358% | 1.618% | 0.9598% | 3.166% | -3.928%
| Z-Score (Absolute) |  | 2.39 | 0 | 0.67 | 0.41 | 2.61 | 0.46 | 0.1 | 0.03 | 2.23 | 5.13 | 0.27 | 1.2 | 0.31 | 0.92 | 0.65 | 1.56 | 1.37
| Analytical results for Sm in OREAS 313 (Certified Value 5.11 ppm) | [22 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 5.4 | 4.7 | 5.645 | 4.7 | 5 | 5 | 5.07 | 4.93 | 5.12 | [12 empty] | 1
| 1 | 2 | 5.4 | 5.7 | 5.4222 | 5.2 | 5.2 | 5 | 5.01 | 4.91 | 5.08 | [12 empty] | 24
| 1 | 3 | 5.6 | 5.1 | 5.4451 | 4 | 5 | 5 | 4.85 | 5.07 | 5 | [12 empty] | 16
| 1 | 4 | 5.5 | 5.7 | 5.3743 | 4.6 | 5 | 5 | 4.8 | 4.86 | 4.96 | [12 empty] | 5.1112
| 1 | 5 | 5.5 | 4.8 | 5.6186 | 4.8 | 5.1 | 5 | 4.91 | 4.95 | 5.15 | [12 empty] | 178
| 1 | 6 | 5.4 | 5.2 | 5.4607 | 4.2 | 5.1 | 4 | 5.21 | 4.78 | 5.11
| Mean |  | 5.4667 | 5.2 | 5.4943 | 4.5833 | 5.0667 | 4.8333 | 4.975 | 4.9167 | 5.07
| Median |  | 5.45 | 5.15 | 5.4529 | 4.65 | 5.05 | 5 | 4.96 | 4.92 | 5.095
| Std Dev. |  | 0.08165 | 0.42895 | 0.11074 | 0.43089 | 0.08165 | 0.40825 | 0.15228 | 0.096678 | 0.074297
| Rel.Std.Dev. |  | 1.494% | 8.249% | 2.016% | 9.401% | 1.612% | 8.447% | 3.061% | 1.966% | 1.465%
| PDM3 |  | 6.955% | 1.738% | 7.497% | -10.33% | -0.8706% | -5.436% | -2.664% | -3.805% | -0.8054%
| Z-Score (Absolute) |  | 1.9 | 0.63 | 2.03 | 2.31 | 0.01 | NA | 0.44 | 0.72 | 0.01
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 313 (Certified Value 2.78 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.17 | 2.18 | 2.19 | 2.21 | 2.24 | [10 empty] | ppm
|  |  | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | BF*OES/MS | PF*OES | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 5 | 3 | 2.8559 | 2.6 | 3 | < 100 | < 50 | 7 | 2 | < 50 | < 5 | [10 empty] | 1
| 1 | 2 | < 5 | 3 | 2.8887 | 2.8 | 3 | < 100 | < 50 | 24 | 2 | < 50 | < 5 | [10 empty] | 31
| 1 | 3 | < 5 | 3 | 2.8317 | 3.1 | 3 | < 100 | < 50 | < 5 | 2 | < 50 | < 5 | [10 empty] | 16
| 1 | 4 | < 5 | 3 | 2.9559 | 2.9 | 3 | < 100 | < 50 | < 5 | 2 | < 50 | 6 | [10 empty] | 2.7796
| 1 | 5 | < 5 | 2 | 2.7867 | 3.1 | 3 | < 100 | < 50 | < 5 | 2 | < 50 | < 5 | [10 empty] | 179
| 1 | 6 | < 5 | 3 | 2.7695 | 2.8 | 2 | < 100 | < 50 | < 5 | 2 | < 50 | < 5
| Mean |  |  | 2.8333 | 2.8481 | 2.8833 | 2.8333 |  |  | 15.5 | 2 |  | 6
| Median |  |  | 3 | 2.8438 | 2.85 | 3 |  |  | 15.5 | 2 |  | 6
| Std Dev. |  |  | 0.40825 | 0.06863 | 0.19408 | 0.40825 |  |  | 12.0208 | 0
| Rel.Std.Dev. |  |  | 14.41% | 2.41% | 6.731% | 14.41% |  |  | 77.55% | 0%
| PDM3 |  |  | 1.933% | 2.463% | 3.731% | 1.933% |  |  | 457.6% | -28.05% |  | 115.9%
| Z-Score (Absolute) |  | 0.67 | 0.09 | 0.06 | 0 | 0.09 | 82.88 | 38.9 | 6.95 | 1.55 | 38.9 | 0.35
| Analytical results for Sr in OREAS 313 (Certified Value 46.6 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 40 | 49.31 | 42.3 | 44.6 | 40.9546 | 54 | 43 | 44.4173 | 52 | 50 | 45.8 | 48 | 42 | 46 | 47 | [6 empty] | 1
| 1 | 2 | 40 | 44.32 | 42.5 | 45.8 | 42.6659 | 56 | 44 | 48.098 | 55 | 50 | 45.4 | 49 | 42 | 47 | 47 | [6 empty] | 32
| 1 | 3 | 40 | 49.46 | 43.3 | 44.9 | 41.0694 | 55 | 44 | 47.6164 | 52 | 50 | 45.9 | 48 | 42 | 44 | 49 | [6 empty] | 16
| 1 | 4 | 40 | 47.14 | 43.7 | 45.1 | 41.9704 | 59 | 46 | 45.1703 | 53 | 50 | 45.8 | 50 | 42 | 45 | 46 | [6 empty] | 46.6095
| 1 | 5 | 40 | 48.82 | 42.4 | 45.2 | 43.4116 | 60 | 45 | 46.5963 | 52 | 50 | 45.1 | 49 | 42 | 46 | 49 | [6 empty] | 180
| 1 | 6 | 40 | 51.4 | 42.7 | 45.5 | 42.2405 | 47 | 44 | 45.5764 | 54 | 50 | 46.1 | 49 | 42 | 46 | 47
| Mean |  | 40 | 48.4083 | 42.8167 | 45.1833 | 42.0521 | 55.1667 | 44.3333 | 46.2458 | 53 | 50 | 45.6833 | 48.8333 | 42 | 45.6667 | 47.5
| Median |  | 40 | 49.065 | 42.6 | 45.15 | 42.1054 | 55.5 | 44 | 46.0863 | 52.5 | 50 | 45.8 | 49 | 42 | 46 | 47
| Std Dev. |  | 0 | 2.4233 | 0.56006 | 0.42622 | 0.94208 | 4.6224 | 1.0328 | 1.4409 | 1.2649 | 0 | 0.3656 | 0.75277 | 0 | 1.0328 | 1.2247
| Rel.Std.Dev. |  | 0% | 5.006% | 1.308% | 0.9433% | 2.24% | 8.379% | 2.33% | 3.116% | 2.387% | 0% | 0.8003% | 1.542% | 0% | 2.262% | 2.578%
| PDM3 |  | -14.18% | 3.859% | -8.137% | -3.06% | -9.778% | 18.36% | -4.883% | -0.7802% | 13.71% | 7.274% | -1.987% | 4.771% | -9.89% | -2.023% | 1.911%
| Z-Score (Absolute) |  | NA | 0.67 | 0.71 | 0.12 | 0.9 | 2.35 | 0.33 | 0.14 | 1.81 | NA | 0 | 0.78 | 0.91 | 0 | 0.45
| Analytical results for Ta in OREAS 313 (Certified Value 0.72 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | 0.7 | 0.7 | 0.7762 | 2 | 0.6 | 0.47166 | 0.7 | 0.4 | 0.7 | 0.7 | < 3 | [10 empty] | 1
| 1 | 2 | 0.8 | 0.7 | 0.76709 | 2 | 0.7 | 0.53557 | 0.6 | 0.9 | 0.6 | 0.9 | < 3 | [10 empty] | 11
| 1 | 3 | 0.7 | 0.7 | 0.78713 | 2.7 | 0.7 | 0.43577 | 0.7 | 0.5 | 0.7 | 0.8 | < 3 | [10 empty] | 16
| 1 | 4 | 0.7 | 0.7 | 0.78918 | 1.9 | 0.8 | 0.50115 | 0.8 | 0.5 | 0.6 | 0.7 | < 3 | [10 empty] | 0.71618
| 1 | 5 | 0.6 | 0.7 | 0.84538 | 1.7 | 0.7 | 0.47499 | 0.7 | 0.4 | 0.6 | 0.7 | < 3 | [10 empty] | 181
| 1 | 6 | 0.7 | 0.8 | 0.81456 | 3.1 | 0.6 | 0.4255 | 0.9 | 0.3 | 0.6 | 0.7 | < 3
| Mean |  | 0.7 | 0.71667 | 0.79659 | 2.2333 | 0.68333 | 0.47411 | 0.73333 | 0.5 | 0.63333 | 0.75
| Median |  | 0.7 | 0.7 | 0.78815 | 2 | 0.7 | 0.47333 | 0.7 | 0.45 | 0.6 | 0.7
| Std Dev. |  | 0.063246 | 0.040825 | 0.028753 | 0.54283 | 0.075277 | 0.040861 | 0.10328 | 0.20976 | 0.05164 | 0.083666
| Rel.Std.Dev. |  | 9.035% | 5.696% | 3.61% | 24.31% | 11.02% | 8.618% | 14.08% | 41.95% | 8.154% | 11.16%
| PDM3 |  | -2.259% | 0.068% | 11.23% | 211.8% | -4.586% | -33.8% | 2.395% | -30.19% | -11.57% | 4.722%
| Z-Score (Absolute) |  | 0.14 | 0 | 0.67 | 12.8 | 0.28 | 2.05 | 0.14 | 1.83 | 0.7 | 0.28 | 6.61
| Analytical results for Tb in OREAS 313 (Certified Value 0.62 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 0.6 | 0.6 | 0.47452 | 0.6 | 0.6 | < 2 | 0.61 | 0.65 | 0.58 | [12 empty] | 1
| 1 | 2 | 0.7 | 0.6 | 0.49414 | 0.6 | 0.61 | < 2 | 0.62 | 0.64 | 0.61 | [12 empty] | 12
| 1 | 3 | 0.7 | 0.6 | 0.461 | 0.7 | 0.61 | < 2 | 0.62 | 0.66 | 0.6 | [12 empty] | 16
| 1 | 4 | 0.7 | 0.7 | 0.50586 | 0.6 | 0.63 | < 2 | 0.6 | 0.63 | 0.6 | [12 empty] | 0.61929
| 1 | 5 | 0.6 | 0.6 | 0.48098 | 0.5 | 0.64 | < 2 | 0.61 | 0.63 | 0.64 | [12 empty] | 182
| 1 | 6 | 0.6 | 0.6 | 0.49072 | 0.5 | 0.67 | < 2 | 0.63 | 0.65 | 0.57
| Mean |  | 0.65 | 0.61667 | 0.48454 | 0.58333 | 0.62667 |  | 0.615 | 0.64333 | 0.6
| Median |  | 0.65 | 0.6 | 0.48585 | 0.6 | 0.62 |  | 0.615 | 0.645 | 0.6
| Std Dev. |  | 0.054772 | 0.040825 | 0.015823 | 0.075277 | 0.02582 |  | 0.010488 | 0.012111 | 0.024495
| Rel.Std.Dev. |  | 8.427% | 6.62% | 3.266% | 12.9% | 4.12% |  | 1.705% | 1.882% | 4.082%
| PDM3 |  | 4.96% | -0.4229% | -21.76% | -5.805% | 1.192% |  | -0.692% | 3.883% | -3.114%
| Z-Score (Absolute) |  | 0.84 | 0 | 3.34 | 0.84 | 0.25 | 9.69 | 0.04 | 0.67 | 0.42
| Analytical results for Te in OREAS 313 (Indicative Value < 1 ppm) | [22 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.12 | 2.13 | 2.14 | 2.24 | [17 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [17 empty] | 2
2
| 1 | 1 | < 6 | < 1 | < 1 | < 5 | [17 empty] | 1
| 1 | 2 | < 6 | < 1 | < 1 | < 5 | [17 empty] | 13
| 1 | 3 | < 6 | < 1 | < 1 | < 5 | [17 empty] | 16
| 1 | 4 | < 6 | < 1 | < 1 | < 5 | [17 empty] | < 1
| 1 | 5 | < 6 | < 1 | < 1 | < 5 | [17 empty] | 32
| 1 | 6 | < 6 | < 1 | < 1 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Th in OREAS 313 (Certified Value 10.7 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.19 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*OES/MS | [10 empty] | 1
2
| 1 | 1 | 10.3 | 10.7 | 10.6 | 11.2028 | 11.2 | 10.7 | 10.5093 | 10.8 | 10.3 | 10.6 | 10 | [10 empty] | 1
| 1 | 2 | 10.5 | 10.7 | 10.6 | 11.1065 | 12 | 11 | 10.5702 | 10.8 | 10.2 | 10.6 | 10 | [10 empty] | 14
| 1 | 3 | 10.4 | 10.6 | 10.5 | 10.8709 | 11.2 | 10.7 | 10.2494 | 10.7 | 10.2 | 10.8 | 10 | [10 empty] | 16
| 1 | 4 | 10.2 | 10.6 | 10.3 | 11.4171 | 11.9 | 10.9 | 10.5914 | 10.7 | 10 | 10.5 | 10 | [10 empty] | 10.6504
| 1 | 5 | 10.2 | 10.8 | 10.6 | 11.5962 | 12 | 10.8 | 10.6759 | 10.6 | 10 | 10.6 | 11 | [10 empty] | 183
| 1 | 6 | 10.3 | 10.6 | 10.7 | 11.1721 | 10.1 | 10.9 | 10.2633 | 10.8 | 10.3 | 10.6 | 10
| Mean |  | 10.3167 | 10.6667 | 10.55 | 11.2276 | 11.4 | 10.8333 | 10.4766 | 10.7333 | 10.1667 | 10.6167 | 10.1667
| Median |  | 10.3 | 10.65 | 10.6 | 11.1874 | 11.55 | 10.85 | 10.5398 | 10.75 | 10.2 | 10.6 | 10
| Std Dev. |  | 0.1169 | 0.08165 | 0.13784 | 0.25194 | 0.74027 | 0.12111 | 0.17883 | 0.08165 | 0.13663 | 0.098319 | 0.40825
| Rel.Std.Dev. |  | 1.133% | 0.7655% | 1.307% | 2.244% | 6.494% | 1.118% | 1.707% | 0.7607% | 1.344% | 0.9261% | 4.016%
| PDM3 |  | -3.133% | 0.1529% | -0.9425% | 5.42% | 7.038% | 1.718% | -1.632% | 0.7789% | -4.542% | -0.3166% | -4.542%
| Z-Score (Absolute) |  | 0.93 | 0.16 | 0.21 | 1.9 | 2.44 | 0.67 | 0.44 | 0.36 | 1.4 | 0 | 1.4
| Analytical results for Ti in OREAS 313 (Certified Value 0.237 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.23 | 0.23129 | 0.2 | 0.24 | 0.2337 | 0.24 | 0.222 | 0.24 | 0.234 | 0.234 | 0.24 | 0.24 | 0.2481 | 0.24 | 0.25 | 0.246 | 0.25 | 0.223 | 0.24 | 0.23 |  | 1
| 1 | 2 | 0.24 | 0.23101 | 0.22 | 0.241 | 0.23353 | 0.234 | 0.228 | 0.24 | 0.228 | 0.234 | 0.24 | 0.23 | 0.24983 | 0.24 | 0.26 | 0.246 | 0.25 | 0.22 | 0.24 | 0.229 |  | 37
| 1 | 3 | 0.23 | 0.23451 | 0.22 | 0.241 | 0.22416 | 0.234 | 0.24 | 0.24 | 0.234 | 0.234 | 0.24 | 0.24 | 0.2481 | 0.25 | 0.25 | 0.246 | 0.24 | 0.223 | 0.24 | 0.232 |  | 16
| 1 | 4 | 0.23 | 0.23552 | 0.21 | 0.241 | 0.23141 | 0.234 | 0.228 | 0.24 | 0.234 | 0.234 | 0.24 | 0.24 | 0.24861 | 0.24 | 0.25 | 0.24 | 0.24 | 0.224 | 0.24 | 0.234 |  | 0.2374
| 1 | 5 | 0.23 | 0.2303 | 0.19 | 0.24 | 0.23282 | 0.24 | 0.228 | 0.222 | 0.234 | 0.24 | 0.24 | 0.24 | 0.24891 | 0.24 | 0.25 | 0.246 | 0.24 | 0.226 | 0.24 | 0.241 |  | 184
| 1 | 6 | 0.23 | 0.23083 | 0.18 | 0.243 | 0.22931 | 0.234 | 0.228 | 0.24 | 0.228 | 0.234 | 0.24 | 0.23 | 0.25012 | 0.24 | 0.25 | 0.246 | 0.25 | 0.221 | 0.24 | 0.238
| Mean |  | 0.23167 | 0.23224 | 0.20333 | 0.241 | 0.23082 | 0.236 | 0.229 | 0.237 | 0.232 | 0.235 | 0.24 | 0.23667 | 0.24895 | 0.24167 | 0.25167 | 0.245 | 0.245 | 0.22283 | 0.24 | 0.234
| Median |  | 0.23 | 0.23115 | 0.205 | 0.241 | 0.23212 | 0.234 | 0.228 | 0.24 | 0.234 | 0.234 | 0.24 | 0.24 | 0.24876 | 0.24 | 0.25 | 0.246 | 0.245 | 0.223 | 0.24 | 0.233
| Std Dev. |  | 0.0040825 | 0.0021943 | 0.01633 | 0.0010954 | 0.0036504 | 0.0030984 | 0.0058992 | 0.0073485 | 0.0030984 | 0.0024495 | 0 | 0.005164 | 0.00086002 | 0.0040825 | 0.0040825 | 0.0024495 | 0.0054772 | 0.002137 | 0 | 0.0046904
| Rel.Std.Dev. |  | 1.762% | 0.9448% | 8.031% | 0.4545% | 1.581% | 1.313% | 2.576% | 3.101% | 1.336% | 1.042% | 0% | 2.182% | 0.3455% | 1.689% | 1.622% | 0.9998% | 2.236% | 0.959% | 0% | 2.004%
| PDM3 |  | -2.416% | -2.172% | -14.35% | 1.516% | -2.771% | -0.5904% | -3.539% | -0.1692% | -2.275% | -1.012% | 1.094% | -0.3096% | 4.863% | 1.797% | 6.009% | 3.201% | 3.201% | -6.137% | 1.094% | -1.433%
| Z-Score (Absolute) |  | 0.66 | 0.58 | 4.75 | 0.69 | 0.78 | 0.06 | 1.07 | 0.08 | 0.64 | 0.21 | 0.54 | 0.06 | 1.84 | 0.79 | 2.23 | 1.24 | 1.27 | 1.94 | 0.54 | 0.32
| Analytical results for Tl in OREAS 313 (Certified Value 20.1 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.04 | 2.12 | 2.13 | 2.14 | 2.16 | 2.23 | 2.24 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [13 empty] | 1
2
| 1 | 1 | 20.9 | 18.6 | 19.9 | 19 | 19.3163 | 20.6 | 20 | 20.02 | [13 empty] | 1
| 1 | 2 | 21 | 19.2 | 20.5 | 19.2 | 20.0038 | 21.4 | 19.9 | 19.86 | [13 empty] | 15
| 1 | 3 | 21.5 | 19.2 | 20.4 | 18.9 | 20.5578 | 21.8 | 20.6 | 20.06 | [13 empty] | 16
| 1 | 4 | 20.7 | 18.7 | 20.4 | 19.1 | 19.2885 | 21.3 | 20.1 | 19.15 | [13 empty] | 20.0867
| 1 | 5 | 21.3 | 19 | 20 | 19.2 | 20.1008 | 20.8 | 20.7 | 21.02 | [13 empty] | 185
| 1 | 6 | 21.2 | 19 | 19.2 | 19.2 | 20.1655 | 21.1 | 20.4 | 20.62
| Mean |  | 21.1 | 18.95 | 20.0667 | 19.1 | 19.9055 | 21.1667 | 20.2833 | 20.1217
| Median |  | 21.1 | 19 | 20.2 | 19.15 | 20.0523 | 21.2 | 20.25 | 20.04
| Std Dev. |  | 0.28983 | 0.251 | 0.48854 | 0.12649 | 0.50372 | 0.43205 | 0.33116 | 0.6453
| Rel.Std.Dev. |  | 1.374% | 1.325% | 2.435% | 0.6623% | 2.531% | 2.041% | 1.633% | 3.207%
| PDM3 |  | 5.045% | -5.659% | -0.09986% | -4.912% | -0.9024% | 5.376% | 0.9788% | 0.174%
| Z-Score (Absolute) |  | 1.15 | 1.3 | 0.03 | 1.13 | 0.22 | 1.22 | 0.22 | 0.03
| Analytical results for Tm in OREAS 313 (Certified Value 0.31 ppm) | [22 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | < 1 | 0.3 | 0.30823 | 0.3 | 0.29 | < 2 | 0.3 | 0.31 | 0.29 | [12 empty] | 1
| 1 | 2 | < 1 | 0.3 | 0.30056 | 0.3 | 0.31 | < 2 | 0.31 | 0.29 | 0.27 | [12 empty] | 8
| 1 | 3 | < 1 | 0.3 | 0.28592 | 0.3 | 0.3 | < 2 | 0.3 | 0.31 | 0.27 | [12 empty] | 16
| 1 | 4 | < 1 | 0.3 | 0.30101 | 0.3 | 0.32 | < 2 | 0.31 | 0.31 | 0.29 | [12 empty] | 0.30513
| 1 | 5 | < 1 | 0.4 | 0.29606 | 0.2 | 0.31 | < 2 | 0.31 | 0.33 | 0.28 | [12 empty] | 186
| 1 | 6 | < 1 | 0.3 | 0.32201 | 0.3 | 0.32 | < 2 | 0.3 | 0.31 | 0.28
| Mean |  |  | 0.31667 | 0.3023 | 0.28333 | 0.30833 |  | 0.305 | 0.31 | 0.28
| Median |  |  | 0.3 | 0.30079 | 0.3 | 0.31 |  | 0.305 | 0.31 | 0.28
| Std Dev. |  |  | 0.040825 | 0.01213 | 0.040825 | 0.01169 |  | 0.0054772 | 0.012649 | 0.0089443
| Rel.Std.Dev. |  |  | 12.89% | 4.013% | 14.41% | 3.791% |  | 1.796% | 4.08% | 3.194%
| PDM3 |  |  | 3.782% | -0.9262% | -7.142% | 1.051% |  | -0.04156% | 1.597% | -8.235%
| Z-Score (Absolute) |  | 15.51 | 0.67 | 0.49 | 2.02 | 0 | 55.97 | 0.27 | 0.13 | 2.29
| Analytical results for U in OREAS 313 (Certified Value 3.61 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.19 | 2.22 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 3.5 | < 0.5 | 3.6 | 3.675 | 3.7 | 3.66 | 3.3794 | 3.7 | 3.6 | 3.49 | 3.6 | 4 | [9 empty] | 1
| 1 | 2 | 3.7 | < 0.5 | 3.8 | 3.6159 | 4 | 3.66 | 3.3513 | 3.7 | 3.6 | 3.56 | 3.8 | 4 | [9 empty] | 17
| 1 | 3 | 3.4 | < 0.5 | 3.6 | 3.3584 | 3.8 | 3.59 | 3.5702 | 4.1 | 3.6 | 3.64 | 3.7 | 4 | [9 empty] | 16
| 1 | 4 | 3.6 | < 0.5 | 3.8 | 3.3821 | 3.6 | 3.62 | 3.6109 | 3.7 | 3.6 | 3.62 | 3.7 | 4 | [9 empty] | 3.6061
| 1 | 5 | 3.5 | < 0.5 | 3.6 | 3.3364 | 4 | 3.65 | 3.4816 | 3.6 | 3.6 | 3.54 | 3.9 | 4 | [9 empty] | 187
| 1 | 6 | 3.4 | < 0.5 | 3.6 | 3.3157 | 3.5 | 3.63 | 3.4406 | 3.5 | 3.6 | 3.65 | 3.7 | 4
| Mean |  | 3.5167 |  | 3.6667 | 3.4473 | 3.7667 | 3.635 | 3.4723 | 3.7167 | 3.6 | 3.5833 | 3.7333 | 4
| Median |  | 3.5 |  | 3.6 | 3.3702 | 3.75 | 3.64 | 3.4611 | 3.7 | 3.6 | 3.59 | 3.7 | 4
| Std Dev. |  | 0.1169 |  | 0.10328 | 0.15624 | 0.20656 | 0.027386 | 0.10309 | 0.20412 | 0 | 0.063456 | 0.10328 | 0
| Rel.Std.Dev. |  | 3.324% |  | 2.817% | 4.532% | 5.484% | 0.7534% | 2.969% | 5.492% | 0% | 1.771% | 2.766% | 0%
| PDM3 |  | -2.481% |  | 1.679% | -4.405% | 4.452% | 0.8007% | -3.711% | 3.065% | -0.1698% | -0.632% | 3.528% | 10.92%
| Z-Score (Absolute) |  | 0.48 | 19.36 | 0.39 | 0.88 | 0.96 | 0.2 | 0.74 | 0.67 | 0 | 0.1 | 0.77 | NA
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 313 (Certified Value 137 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 130 | 137.1415 | 142.8 | 142 | 150 | 139 | 141.6288 | 138 | 140 | 155 | 133 | 128 | 138 | 131 | [7 empty] | 1
| 1 | 2 | 140 | 133.8165 | 142.2 | 138 | 146 | 138 | 140.0901 | 138 | 140 | 162 | 135 | 127 | 140 | 132 | [7 empty] | 18
| 1 | 3 | 140 | 138.4905 | 141.2 | 144 | 145 | 139 | 137.971 | 139 | 140 | 157 | 133 | 127 | 137 | 131 | [7 empty] | 16
| 1 | 4 | 140 | 137.9585 | 142.6 | 139 | 144 | 136 | 134.3355 | 141 | 140 | 156 | 134 | 127 | 141 | 127 | [7 empty] | 137.148
| 1 | 5 | 130 | 138.2435 | 144.1 | 139 | 145 | 137 | 137.2102 | 143 | 140 | 161 | 130 | 129 | 136 | 136 | [7 empty] | 188
| 1 | 6 | 130 | 134.3235 | 138.3 | 141 | 142 | 139 | 133.2457 | 144 | 140 | 155 | 135 | 127 | 139 | 130
| Mean |  | 135 | 136.6623 | 141.8667 | 140.5 | 145.3333 | 138 | 137.4135 | 140.5 | 140 | 157.6667 | 133.3333 | 127.5 | 138.5 | 131.1667
| Median |  | 135 | 137.55 | 142.4 | 140 | 145 | 138.5 | 137.5906 | 140 | 140 | 156.5 | 133.5 | 127 | 138.5 | 131
| Std Dev. |  | 5.4772 | 2.065 | 1.9836 | 2.2583 | 2.6583 | 1.2649 | 3.2289 | 2.5884 | 0 | 3.0768 | 1.8619 | 0.83666 | 1.8708 | 2.9269
| Rel.Std.Dev. |  | 4.057% | 1.511% | 1.398% | 1.607% | 1.829% | 0.9166% | 2.35% | 1.842% | 0% | 1.951% | 1.396% | 0.6562% | 1.351% | 2.231%
| PDM3 |  | -1.566% | -0.3541% | 3.441% | 2.444% | 5.968% | 0.6212% | 0.1936% | 2.444% | 2.08% | 14.96% | -2.781% | -7.035% | 0.9858% | -4.361%
| Z-Score (Absolute) |  | 0.67 | 0.3 | 0.87 | 0.56 | 1.65 | 0 | 0.13 | 0.56 | NA | 4.42 | 1.05 | 2.36 | 0.11 | 1.54
Results from laboratory 2.17 were removed due to their 10 ppm reading resolution.
| Analytical results for W in OREAS 313 (Indicative Value 2.02 ppm) | [22 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [8 empty] | ppm
|  |  | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | < 5 | 2 | < 5 | 2.6 | 2 | 1.3103 | 3 | < 100 | 1.7 | 508 | < 50 | 3 | < 5 | [8 empty] | 1
| 1 | 2 | < 5 | 2 | < 5 | 1.9 | 2 | 1.3129 | 2 | < 100 | 1.9 | 509 | < 50 | 2 | < 5 | [8 empty] | 19
| 1 | 3 | < 5 | 1 | < 5 | 2.5 | 2 | 1.3643 | 3 | < 100 | 1.8 | 502 | < 50 | 2 | < 5 | [8 empty] | 16
| 1 | 4 | < 5 | 2 | < 5 | 3 | 2 | 1.2943 | 5 | < 100 | 1.6 | 499 | < 50 | 3 | < 5 | [8 empty] | 2.022
| 1 | 5 | < 5 | 2 | < 5 | 2.4 | 2 | 1.3683 | 2 | < 100 | 1.8 | 495 | < 50 | 2 | < 5 | [8 empty] | 33
| 1 | 6 | < 5 | 2 | < 5 | 2.4 | 2 | 1.3748 | 2 | < 100 | 1.9 | 496 | < 50 | 2 | < 5
| Mean |  |  | 1.8333 |  | 2.4667 | 2 | 1.3375 | 2.8333 |  | 1.7833 | 501.5 |  | 2.3333
| Median |  |  | 2 |  | 2.45 | 2 | 1.3386 | 2.5 |  | 1.8 | 500.5 |  | 2
| Std Dev. |  |  | 0.40825 |  | 0.3559 | 0 | 0.035413 | 1.169 |  | 0.1169 | 5.9582 |  | 0.5164
| Rel.Std.Dev. |  |  | 22.27% |  | 14.43% | 0% | 2.648% | 41.26% |  | 6.555% | 1.188% |  | 22.13%
| PDM3 |  |  | -9.332% |  | 21.99% | -1.089% | -33.85% | 40.12% |  | -11.8% | 24702% |  | 15.4%
| Z-Score (Absolute) |  | 0 | 0.9 | 0 | 0.04 | 0.67 | 1.57 | 0.45 | 64.06 | 0.97 | 672.96 | 30.34 | 0.22 | 0
| Analytical results for Y in OREAS 313 (Certified Value 20.1 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*OES | PF*OES | BF*MS | PF*MS | PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 20 | 18.9 | 20 | 17.3731 | 19.8 | 20.4 | 20.3096 | 19.5 | 18 | 20.9 | 21 | 16 | 20.4 | 21 | 19 | [6 empty] | 1
| 1 | 2 | 20.1 | 20 | 20.2 | 18.2948 | 22.5 | 20.9 | 19.8742 | 20.6 | 20 | 20.8 | 20 | 16 | 20.1 | 21.4 | 19 | [6 empty] | 20
| 1 | 3 | 20 | 19.6 | 20.5 | 17.8979 | 19.8 | 20.6 | 20.1256 | 20.7 | 21 | 20.7 | 21 | 17 | 20.35 | 20.5 | 19 | [6 empty] | 16
| 1 | 4 | 19.8 | 19.2 | 20.7 | 18.8475 | 20.7 | 21.4 | 19.8063 | 20.4 | 21 | 20.9 | 20 | 16 | 20.2 | 20.8 | 19 | [6 empty] | 20.0938
| 1 | 5 | 20 | 17.4 | 20.4 | 18.2014 | 20.5 | 20.5 | 20.1291 | 20.3 | 20 | 20.2 | 20 | 16 | 19.9 | 19.9 | 19 | [6 empty] | 189
| 1 | 6 | 19.9 | 17.7 | 20.7 | 18.0973 | 19.8 | 20.8 | 19.7686 | 19.7 | 18 | 20.8 | 20 | 17 | 20.05 | 21.1 | 19
| Mean |  | 19.9667 | 18.8 | 20.4167 | 18.1187 | 20.5167 | 20.7667 | 20.0022 | 20.2 | 19.6667 | 20.7167 | 20.3333 | 16.3333 | 20.1667 | 20.7833 | 19
| Median |  | 20 | 19.05 | 20.45 | 18.1494 | 20.15 | 20.7 | 19.9999 | 20.35 | 20 | 20.8 | 20 | 16 | 20.15 | 20.9 | 19
| Std Dev. |  | 0.10328 | 1.0412 | 0.27869 | 0.48458 | 1.0496 | 0.36148 | 0.21687 | 0.4899 | 1.3663 | 0.26394 | 0.5164 | 0.5164 | 0.18886 | 0.52694 | 0
| Rel.Std.Dev. |  | 0.5173% | 5.538% | 1.365% | 2.674% | 5.116% | 1.741% | 1.084% | 2.425% | 6.947% | 1.274% | 2.54% | 3.162% | 0.9365% | 2.535% | 0%
| PDM3 |  | -0.6325% | -6.439% | 1.607% | -9.829% | 2.105% | 3.349% | -0.4555% | 0.5287% | -2.126% | 3.1% | 1.192% | -18.71% | 0.3628% | 3.432% | -5.443%
| Z-Score (Absolute) |  | 0.27 | 1.84 | 0.34 | 2.76 | 0.47 | 0.81 | 0.22 | 0.04 | 0.67 | 0.74 | 0.22 | 5.17 | 0 | 0.83 | 1.57
| Analytical results for Yb in OREAS 313 (Certified Value 2.06 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 1.9 | 2 | 1.9595 | 2 | 1.9 | < 2 | 2.07 | 2.05 | 2.02 | [12 empty] | 1
| 1 | 2 | 2 | 1.92 | 1.9979 | 1.7 | 2.1 | 2 | 2.14 | 2.15 | 2 | [12 empty] | 21
| 1 | 3 | 1.9 | 2.07 | 2.1355 | 2.3 | 2 | 2 | 2.01 | 2.05 | 2.09 | [12 empty] | 16
| 1 | 4 | 2 | 2.11 | 2.1235 | 2.7 | 2.1 | 2 | 2.06 | 1.9 | 1.99 | [12 empty] | 2.0567
| 1 | 5 | 1.7 | 2.14 | 2.1917 | 2.1 | 2.2 | < 2 | 2.1 | 2.05 | 2.07 | [12 empty] | 190
| 1 | 6 | 1.9 | 2.06 | 2.0615 | 2.2 | 2.1 | < 2 | 2.12 | 1.95 | 2.05
| Mean |  | 1.9 | 2.05 | 2.0783 | 2.1667 | 2.0667 | 2 | 2.0833 | 2.025 | 2.0367
| Median |  | 1.9 | 2.065 | 2.0925 | 2.15 | 2.1 | 2 | 2.085 | 2.05 | 2.035
| Std Dev. |  | 0.10954 | 0.079498 | 0.088346 | 0.33267 | 0.10328 | 0 | 0.046762 | 0.088034 | 0.039833
| Rel.Std.Dev. |  | 5.766% | 3.878% | 4.251% | 15.35% | 4.997% | 0% | 2.245% | 4.347% | 1.956%
| PDM3 |  | -7.617% | -0.3236% | 1.051% | 5.349% | 0.4868% | -2.755% | 1.297% | -1.539% | -0.9719%
| Z-Score (Absolute) |  | 4.58 | 0.24 | 0.58 | 3.13 | 0.24 | NA | 0.72 | 0.96 | 0.63
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 313 (Certified Value 3.45 wt.%) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 2
3
| 1 | 1 | 3.29 | 3.3367 | > 1 | 3.3746 | 3.48 | 3.27 | 3.5 | 3.4 | 3.46 | > 1 | 3.48 | 3.5616 | 3.5056 | 3.448 | 3.51 | 3.5555 | 3.3981 | > 2.5 |  |  |  | 1
| 1 | 2 | 3.29 | 3.3491 | > 1 | 3.3302 | 3.43 | 3.44 | 3.49 | 3.37 | 3.53 | > 1 | 3.42 | 3.619 | 3.4782 | 3.581 | 3.48 | 3.5674 | 3.3714 | > 2.5 |  |  |  | #N/A
| 1 | 3 | 3.29 | 3.3382 | > 1 | 3.3463 | 3.5 | 3.52 | 3.42 | 3.42 | 3.54 | > 1 | 3.43 | 3.577 | 3.5093 | 3.461 | 3.61 | 3.5483 | 3.3822 | > 2.5 |  |  |  | 16
| 1 | 4 | 3.26 | 3.3413 | > 1 | 3.3599 | 3.51 | 3.4 | 3.47 | 3.39 | 3.44 | > 1 | 3.47 | 3.5551 | 3.5157 | 3.49 | 3.44 | 3.5319 | 3.3768 | > 2.5 |  |  |  | 3.4463
| 1 | 5 | 3.27 | 3.362 | > 1 | 3.3701 | 3.47 | 3.35 | 3.43 | 3.39 | 3.59 | > 1 | 3.47 | 3.6121 | 3.4681 | 3.532 | 3.5 | 3.5627 | 3.363 | > 2.5 |  |  |  | 191
| 1 | 6 | 3.25 | 3.3209 | > 1 | 3.379 | 3.44 | 3.47 | 3.49 | 3.41 | 3.48 | > 1 | 3.42 | 3.6085 | 3.5343 | 3.414 | 3.53 | 3.543 | 3.3799 | > 2.5
| Mean |  | 3.275 | 3.3414 |  | 3.36 | 3.4717 | 3.4083 | 3.4667 | 3.3967 | 3.5067 |  | 3.4483 | 3.5889 | 3.5019 | 3.4877 | 3.5117 | 3.5515 | 3.3786
| Median |  | 3.28 | 3.3397 |  | 3.365 | 3.475 | 3.42 | 3.48 | 3.395 | 3.505 |  | 3.45 | 3.5928 | 3.5074 | 3.4755 | 3.505 | 3.5519 | 3.3784
| Std Dev. |  | 0.017607 | 0.013684 |  | 0.018714 | 0.031885 | 0.089312 | 0.033862 | 0.017512 | 0.056451 |  | 0.027869 | 0.027765 | 0.024538 | 0.060652 | 0.057067 | 0.013122 | 0.011776
| Rel.Std.Dev. |  | 0.5376% | 0.4095% |  | 0.557% | 0.9184% | 2.62% | 0.9768% | 0.5156% | 1.61% |  | 0.8082% | 0.7736% | 0.7007% | 1.739% | 1.625% | 0.3695% | 0.3485%
| PDM3 |  | -4.971% | -3.045% |  | -2.504% | 0.7354% | -1.102% | 0.5903% | -1.441% | 1.751% |  | 0.05832% | 4.137% | 1.611% | 1.2% | 1.896% | 3.051% | -1.966%
| Z-Score (Absolute) |  | 2.22 | 1.45 | NA | 1.23 | 0.06 | 0.67 | 0 | 0.81 | 0.46 | NA | 0.21 | 1.41 | 0.41 | 0.24 | 0.52 | 0.98 | 1.02 | NA
| Analytical results for Zr in OREAS 313 (Certified Value 103 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.13 | 2.18 | 2.19 | 2.20 | [14 empty] | ppm
|  |  | PF*OES | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*OES | [14 empty] | 0
0
| 1 | 1 | 78.9 | 107 | 99.8761 | 95.3 | 75 | 111 | 104 | [14 empty] | 1
| 1 | 2 | 79.5 | 105 | 104.5609 | 107 | 98 | 107.5 | 104 | [14 empty] | 22
| 1 | 3 | 81.8 | 99 | 99.5627 | 99.7 | 102 | 110.9 | 109 | [14 empty] | 16
| 1 | 4 | 82.8 | 102 | 102.1105 | 99.4 | 107 | 110.5 | 100 | [14 empty] | 103.3971
| 1 | 5 | 80.7 | 104 | 101.6988 | 103 | 93 | 104.2 | 103 | [14 empty] | 192
| 1 | 6 | 80.7 | 104 | 102.6045 | 91.8 | 73 | 107.2 | 103
| Mean |  | 80.7333 | 103.5 | 101.7356 | 99.3667 | 91.3333 | 108.55 | 103.8333
| Median |  | 80.7 | 104 | 101.9047 | 99.55 | 95.5 | 109 | 103.5
| Std Dev. |  | 1.4348 | 2.7386 | 1.8471 | 5.3921 | 14.208 | 2.7267 | 2.9269
| Rel.Std.Dev. |  | 1.777% | 2.646% | 1.816% | 5.426% | 15.56% | 2.512% | 2.819%
| PDM3 |  | -21.92% | 0.09951% | -1.607% | -3.898% | -11.67% | 4.984% | 0.4219%
| Z-Score (Absolute) |  | 5.98 | 0.5 | 0 | 0.67 | 2.96 | 1.94 | 0.6
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C965: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 313 (Certified Value 1.59 wt.%) | [24 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 1.63 | 1.61 | 1.48 | 1.56 | 1.5731 | 1.54 | 1.66 | 1.59 | 1.59 | 1.59 | 1.57 | 1.598 | 1.4941 | 1.61 | 1.58 | 1.67 | 1.6 | 1.57 | 1.53 | 1.62 | 1.5 | 1.58 | 1.68 | 1
| 1 | 2 | 1.62 | 1.62 | 1.51 | 1.54 | 1.5797 | 1.54 | 1.64 | 1.59 | 1.59 | 1.58 | 1.57 | 1.633 | 1.51 | 1.6 | 1.57 | 1.66 | 1.58 | 1.56 | 1.54 | 1.62 | 1.5 | 1.6 | 1.68 | #N/A
| 1 | 3 | 1.63 | 1.59 | 1.5 | 1.55 | 1.5905 | 1.56 | 1.72 | 1.58 | 1.59 | 1.58 | 1.56 | 1.631 | 1.5 | 1.51 | 1.57 | 1.67 | 1.58 | 1.58 | 1.54 | 1.6 | 1.52 | 1.59 | 1.68 | 16
| 1 | 4 | 1.64 | 1.61 | 1.49 | 1.53 | 1.5672 | 1.53 | 1.67 | 1.59 | 1.6 | 1.58 | 1.57 | 1.622 | 1.5171 | 1.26 | 1.58 | 1.65 | 1.6 | 1.56 | 1.55 | 1.65 | 1.49 | 1.59 | 1.69 | 1.5855
| 1 | 5 | 1.63 | 1.59 | 1.51 | 1.55 | 1.5771 | 1.54 | 1.72 | 1.58 | 1.6 | 1.59 | 1.57 | 1.619 | 1.4965 | 1.61 | 1.57 | 1.67 | 1.62 | 1.58 | 1.56 | 1.6 | 1.51 | 1.6 | 1.68 | 194
| 1 | 6 | 1.63 | 1.61 | 1.49 | 1.53 | 1.5841 | 1.58 | 1.68 | 1.59 | 1.59 | 1.58 | 1.57 | 1.627 | 1.52 | 1.59 | 1.58 | 1.67 | 1.59 | 1.56 | 1.56 | 1.59 | 1.49 | 1.6 | 1.68
| Mean |  | 1.63 | 1.605 | 1.4967 | 1.5433 | 1.5786 | 1.5483 | 1.6817 | 1.5867 | 1.5933 | 1.5833 | 1.5683 | 1.6217 | 1.5063 | 1.53 | 1.575 | 1.665 | 1.595 | 1.5683 | 1.5467 | 1.6133 | 1.5017 | 1.5933 | 1.6817
| Median |  | 1.63 | 1.61 | 1.495 | 1.545 | 1.5784 | 1.54 | 1.675 | 1.59 | 1.59 | 1.58 | 1.57 | 1.6245 | 1.505 | 1.595 | 1.575 | 1.67 | 1.595 | 1.565 | 1.545 | 1.61 | 1.5 | 1.595 | 1.68
| Std Dev. |  | 0.0063246 | 0.012247 | 0.012111 | 0.012111 | 0.0082046 | 0.018348 | 0.032506 | 0.005164 | 0.005164 | 0.005164 | 0.0040825 | 0.012738 | 0.010968 | 0.13755 | 0.0054772 | 0.0083666 | 0.015166 | 0.0098319 | 0.012111 | 0.021602 | 0.01169 | 0.008165 | 0.0040825
| Rel.Std.Dev. |  | 0.388% | 0.7631% | 0.8092% | 0.7847% | 0.5197% | 1.185% | 1.933% | 0.3255% | 0.3241% | 0.3261% | 0.2603% | 0.7855% | 0.7281% | 8.99% | 0.3478% | 0.5025% | 0.9508% | 0.6269% | 0.783% | 1.339% | 0.7785% | 0.5124% | 0.2428%
| PDM3 |  | 2.805% | 1.228% | -5.605% | -2.661% | -0.437% | -2.346% | 6.063% | 0.07163% | 0.4921% | -0.1386% | -1.085% | 2.279% | -4.999% | -3.502% | -0.6642% | 5.012% | 0.5972% | -1.085% | -2.451% | 1.754% | -5.289% | 0.4921% | 6.063%
| Z-Score (Absolute) |  | 0.9 | 0.42 | 1.67 | 0.77 | 0.09 | 0.67 | 1.89 | 0.06 | 0.19 | 0 | 0.29 | 0.74 | 1.48 | 1.03 | 0.16 | 1.57 | 0.22 | 0.29 | 0.71 | 0.58 | 1.57 | 0.19 | 1.89
| Analytical results for S in OREAS 313 (Certified Value 4.66 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 4.6 | 3.34 | 4.52 | 4.62 | 4.6921 | 4.48 | 4.68 | 4.78 | 4.63 | 4.83 | 4.75 | 4.503 | 4.7081 | 4.96 | 4.76 | 4.76 | 4.73 | 4.71 | 4.65 | 4.52 | 4.57 | 4.66 | 4.59 | 1
| 1 | 2 | 4.59 | 3.44 | 4.56 | 4.54 | 4.7033 | 4.78 | 4.7 | 4.72 | 4.67 | 4.66 | 4.75 | 4.567 | 4.7134 | 5 | 4.71 | 4.82 | 4.58 | 4.6 | 4.65 | 4.56 | 4.53 | 4.66 | 4.6 | #N/A
| 1 | 3 | 4.6 | 3.4 | 4.54 | 4.57 | 4.7082 | 4.53 | 4.72 | 4.76 | 4.65 | 4.83 | 4.71 | 4.521 | 4.7527 | 4.82 | 4.73 | 4.82 | 4.71 | 4.72 | 4.66 | 4.57 | 4.53 | 4.66 | 4.59 | 16
| 1 | 4 | 4.58 | 3.5 | 4.54 | 4.51 | 4.6852 | 4.43 | 4.74 | 4.79 | 4.71 | 4.75 | 4.74 | 4.434 | 4.7615 | 4.88 | 4.75 | 4.84 | 4.59 | 4.6 | 4.63 | 4.69 | 4.55 | 4.67 | 4.55 | 4.6628
| 1 | 5 | 4.59 | 3.43 | 4.51 | 4.59 | 4.6761 | 4.72 | 4.74 | 4.74 | 4.68 | 4.77 | 4.79 | 4.487 | 4.7216 | 5.03 | 4.74 | 4.83 | 4.53 | 4.69 | 4.66 | 4.56 | 4.53 | 4.65 | 4.6 | 195
| 1 | 6 | 4.6 | 3.31 | 4.51 | 4.55 | 4.7456 | 4.64 | 4.69 | 4.77 | 4.65 | 4.79 | 4.73 | 4.498 | 4.7196 | 4.86 | 4.73 | 4.74 | 4.47 | 4.7 | 4.68 | 4.62 | 4.54 | 4.63 | 4.57
| Mean |  | 4.5933 | 3.4033 | 4.53 | 4.5633 | 4.7018 | 4.5967 | 4.7117 | 4.76 | 4.665 | 4.7717 | 4.745 | 4.5017 | 4.7295 | 4.925 | 4.7367 | 4.8017 | 4.6017 | 4.67 | 4.655 | 4.5867 | 4.5417 | 4.655 | 4.5833
| Median |  | 4.595 | 3.415 | 4.53 | 4.56 | 4.6977 | 4.585 | 4.71 | 4.765 | 4.66 | 4.78 | 4.745 | 4.5005 | 4.7206 | 4.92 | 4.735 | 4.82 | 4.585 | 4.695 | 4.655 | 4.565 | 4.535 | 4.66 | 4.59
| Std Dev. |  | 0.008165 | 0.069474 | 0.02 | 0.038816 | 0.024468 | 0.13895 | 0.025626 | 0.026077 | 0.028107 | 0.063377 | 0.026646 | 0.043477 | 0.022097 | 0.083845 | 0.017512 | 0.041191 | 0.10128 | 0.055136 | 0.016432 | 0.059889 | 0.016021 | 0.013784 | 0.019664
| Rel.Std.Dev. |  | 0.1778% | 2.041% | 0.4415% | 0.8506% | 0.5204% | 3.023% | 0.5439% | 0.5478% | 0.6025% | 1.328% | 0.5616% | 0.9658% | 0.4672% | 1.702% | 0.3697% | 0.8578% | 2.201% | 1.181% | 0.353% | 1.306% | 0.3528% | 0.2961% | 0.429%
| PDM3 |  | -1.489% | -27.01% | -2.848% | -2.133% | 0.8358% | -1.418% | 1.048% | 2.085% | 0.04753% | 2.335% | 1.763% | -3.455% | 1.431% | 5.624% | 1.585% | 2.979% | -1.311% | 0.1548% | -0.1669% | -1.632% | -2.598% | -0.1669% | -1.704%
| Z-Score (Absolute) |  | 0.56 | 11.33 | 1.13 | 0.83 | 0.42 | 0.53 | 0.51 | 0.95 | 0.09 | 1.06 | 0.81 | 1.39 | 0.67 | 2.44 | 0.74 | 1.33 | 0.48 | 0.14 | 0 | 0.62 | 1.03 | 0 | 0.65
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 313 (Indicative Value 70.3 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 74.4 | 1
| 1 | 2 | 66.2 | 10
| Mean |  | 70.3 | 16
| Median |  | 70.3 | 70.3
| Std Dev. |  | 5.7983 | 36
| Rel.Std.Dev. |  | 8.248%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 313 (Indicative Value 94 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 93.8 | 1
| 1 | 2 | 93.4 | 11
| Mean |  | 93.6 | 16
| Median |  | 93.6 | 93.6
| Std Dev. |  | 0.28284 | 37
| Rel.Std.Dev. |  | 0.3022%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 313 (Indicative Value 0.236 wt.%) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.236 | 1
| 1 | 2 | 0.236 | 2
| Mean |  | 0.236 | 16
| Median |  | 0.236 | 0.236
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 313 (Indicative Value 2.4 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.4 | 33
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 313 (Indicative Value 0.4 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.42 | 1
| 1 | 2 | 0.38 | 34
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0.028284 | 40
| Rel.Std.Dev. |  | 7.071%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 313 (Indicative Value 57 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 56.9 | 1
| 1 | 2 | 57.3 | 14
| Mean |  | 57.1 | 16
| Median |  | 57.1 | 57.1
| Std Dev. |  | 0.28284 | 41
| Rel.Std.Dev. |  | 0.4953%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 313 (Indicative Value 25.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25.2 | 1
| 1 | 2 | 25 | 16
| Mean |  | 25.1 | 16
| Median |  | 25.1 | 25.1
| Std Dev. |  | 0.14142 | 42
| Rel.Std.Dev. |  | 0.5634%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 313 (Indicative Value 71 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72 | 1
| 1 | 2 | 70 | 37
| Mean |  | 71 | 16
| Median |  | 71 | 71
| Std Dev. |  | 1.4142 | 43
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 313 (Indicative Value 3.09 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.14 | 1
| 1 | 2 | 3.03 | 17
| Mean |  | 3.085 | 16
| Median |  | 3.085 | 3.085
| Std Dev. |  | 0.077782 | 44
| Rel.Std.Dev. |  | 2.521%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 313 (Indicative Value 0.036 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.036 | 1
| 1 | 2 | 0.036 | 39
| Mean |  | 0.036 | 16
| Median |  | 0.036 | 0.036
| Std Dev. |  | 0 | 45
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 313 (Indicative Value 3.62 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.68 | 1
| 1 | 2 | 3.56 | 5
| Mean |  | 3.62 | 16
| Median |  | 3.62 | 3.62
| Std Dev. |  | 0.084853 | 46
| Rel.Std.Dev. |  | 2.344%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 313 (Indicative Value 2.15 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.19 | 1
| 1 | 2 | 2.11 | 6
| Mean |  | 2.15 | 16
| Median |  | 2.15 | 2.15
| Std Dev. |  | 0.056569 | 47
| Rel.Std.Dev. |  | 2.631%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 313 (Indicative Value 0.95 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.95 | 7
| Mean |  | 0.95 | 16
| Median |  | 0.95 | 0.95
| Std Dev. |  | 0 | 48
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 313 (Indicative Value 13.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.7 | 1
| 1 | 2 | 13.9 | 21
| Mean |  | 13.8 | 16
| Median |  | 13.8 | 13.8
| Std Dev. |  | 0.14142 | 49
| Rel.Std.Dev. |  | 1.025%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 313 (Indicative Value 3.92 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 4.04 | 8
| Mean |  | 3.92 | 16
| Median |  | 3.92 | 3.92
| Std Dev. |  | 0.16971 | 50
| Rel.Std.Dev. |  | 4.329%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 313 (Indicative Value 8.53 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.5 | 1
| 1 | 2 | 8.55 | 23
| Mean |  | 8.525 | 16
| Median |  | 8.525 | 8.525
| Std Dev. |  | 0.035355 | 51
| Rel.Std.Dev. |  | 0.4147%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 313 (Indicative Value 2.77 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.77 | 1
| 1 | 2 | 2.77 | 30
| Mean |  | 2.77 | 16
| Median |  | 2.77 | 2.77
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 313 (Indicative Value 0.72 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.71 | 1
| 1 | 2 | 0.72 | 4
| Mean |  | 0.715 | 16
| Median |  | 0.715 | 0.715
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 0.989%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 313 (Indicative Value 0.25 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.25 | 11
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 313 (Indicative Value 31.1 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.9 | 1
| 1 | 2 | 31.2 | 12
| Mean |  | 31.05 | 16
| Median |  | 31.05 | 31.05
| Std Dev. |  | 0.21213 | 39
| Rel.Std.Dev. |  | 0.6832%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 313 (Indicative Value 0.28 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.27 | 1
| 1 | 2 | 0.28 | 5
| Mean |  | 0.275 | 16
| Median |  | 0.275 | 0.275
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 313 (Indicative Value 0.183 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.182 | 1
| 1 | 2 | 0.183 | 35
| Mean |  | 0.1825 | 16
| Median |  | 0.1825 | 0.1825
| Std Dev. |  | 0.00070711 | 41
| Rel.Std.Dev. |  | 0.3875%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 313 (Indicative Value 5.4 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.4 | 1
| 1 | 2 | 5.4 | 14
| Mean |  | 5.4 | 16
| Median |  | 5.4 | 5.4
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 313 (Indicative Value 8.29 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.21 | 1
| 1 | 2 | 8.36 | 16
| Mean |  | 8.285 | 16
| Median |  | 8.285 | 8.285
| Std Dev. |  | 0.10607 | 43
| Rel.Std.Dev. |  | 1.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 313 (Indicative Value 27.2 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27.3 | 1
| 1 | 2 | 27.1 | 7
| Mean |  | 27.2 | 16
| Median |  | 27.2 | 27.2
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 0.5199%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 313 (Indicative Value 15 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16 | 1
| 1 | 2 | 14 | 18
| Mean |  | 15 | 16
| Median |  | 15 | 15
| Std Dev. |  | 1.4142 | 45
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 313 (Indicative Value 3.87 wt.%) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.9 | 1
| 1 | 2 | 3.83 | 40
| Mean |  | 3.865 | 16
| Median |  | 3.865 | 3.865
| Std Dev. |  | 0.049497 | 46
| Rel.Std.Dev. |  | 1.281%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 313 (Indicative Value 7.36 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.24 | 1
| 1 | 2 | 7.48 | 19
| Mean |  | 7.36 | 16
| Median |  | 7.36 | 7.36
| Std Dev. |  | 0.16971 | 47
| Rel.Std.Dev. |  | 2.306%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 313 (Indicative Value 106 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 105 | 1
| 1 | 2 | 106 | 20
| Mean |  | 105.5 | 16
| Median |  | 105.5 | 105.5
| Std Dev. |  | 0.70711 | 48
| Rel.Std.Dev. |  | 0.6702%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 313 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 27
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 49
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 313 (Indicative Value 49.5 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 49.3 | 1
| 1 | 2 | 49.7 | 21
| Mean |  | 49.5 | 16
| Median |  | 49.5 | 49.5
| Std Dev. |  | 0.28284 | 50
| Rel.Std.Dev. |  | 0.5714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 313 (Indicative Value 8.45 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.4 | 1
| 1 | 2 | 8.5 | 22
| Mean |  | 8.45 | 16
| Median |  | 8.45 | 8.45
| Std Dev. |  | 0.070711 | 51
| Rel.Std.Dev. |  | 0.8368%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 313 (Indicative Value 5.04 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.08 | 1
| 1 | 2 | 4.99 | 24
| Mean |  | 5.035 | 16
| Median |  | 5.035 | 5.035
| Std Dev. |  | 0.06364 | 36
| Rel.Std.Dev. |  | 1.264%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 313 (Indicative Value 2.6 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.8 | 31
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0.28284 | 37
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 313 (Indicative Value 44.2 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 44.1 | 1
| 1 | 2 | 44.2 | 32
| Mean |  | 44.15 | 16
| Median |  | 44.15 | 44.15
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.1602%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 313 (Indicative Value 0.33 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.33 | 1
| 1 | 2 | 0.32 | 11
| Mean |  | 0.325 | 16
| Median |  | 0.325 | 0.325
| Std Dev. |  | 0.0070711 | 39
| Rel.Std.Dev. |  | 2.176%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 313 (Indicative Value 0.61 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.62 | 12
| Mean |  | 0.605 | 16
| Median |  | 0.605 | 0.605
| Std Dev. |  | 0.021213 | 40
| Rel.Std.Dev. |  | 3.506%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 313 (Indicative Value 0.15 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | < 0.2 | 13
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 41
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 313 (Indicative Value 10.3 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.1 | 1
| 1 | 2 | 10.4 | 14
| Mean |  | 10.25 | 16
| Median |  | 10.25 | 10.25
| Std Dev. |  | 0.21213 | 42
| Rel.Std.Dev. |  | 2.07%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 313 (Indicative Value 0.242 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.24 | 1
| 1 | 2 | 0.243 | 37
| Mean |  | 0.2415 | 16
| Median |  | 0.2415 | 0.2415
| Std Dev. |  | 0.0021213 | 43
| Rel.Std.Dev. |  | 0.8784%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 313 (Indicative Value 0.31 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.31 | 8
| Mean |  | 0.305 | 16
| Median |  | 0.305 | 0.305
| Std Dev. |  | 0.0070711 | 44
| Rel.Std.Dev. |  | 2.318%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 313 (Indicative Value 3.48 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.51 | 1
| 1 | 2 | 3.44 | 17
| Mean |  | 3.475 | 16
| Median |  | 3.475 | 3.475
| Std Dev. |  | 0.049497 | 45
| Rel.Std.Dev. |  | 1.424%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 313 (Indicative Value 137 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 136 | 1
| 1 | 2 | 137 | 18
| Mean |  | 136.5 | 16
| Median |  | 136.5 | 136.5
| Std Dev. |  | 0.70711 | 46
| Rel.Std.Dev. |  | 0.518%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 313 (Indicative Value 0.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.5 | 19
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 313 (Indicative Value 20.3 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20.5 | 20
| Mean |  | 20.25 | 16
| Median |  | 20.25 | 20.25
| Std Dev. |  | 0.35355 | 48
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 313 (Indicative Value 2.11 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.19 | 1
| 1 | 2 | 2.02 | 21
| Mean |  | 2.105 | 16
| Median |  | 2.105 | 2.105
| Std Dev. |  | 0.12021 | 49
| Rel.Std.Dev. |  | 5.711%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 313 (Indicative Value 102 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 101 | 1
| 1 | 2 | 102 | 22
| Mean |  | 101.5 | 16
| Median |  | 101.5 | 101.5
| Std Dev. |  | 0.70711 | 50
| Rel.Std.Dev. |  | 0.6967%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.