## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 353c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Umpire Labs (dry sample basis)
Classical Wet Chemistry
Pb, Lead (wt.%) | 72.2872 | 72.2023 | 72.3721 | 72.2053 | 72.3691
Pb Fire Assay
Ag, Silver (wt.%) | 0.26856 | 0.26731 | 0.26981 | 0.26721 | 0.26992
Geoanalytical Labs ('as received' sample basis)
4-Acid Digestion
Ag, Silver (wt.%) | 0.26836 | 0.25885 | 0.27788 | 0.26322 | 0.2735
Al, Aluminium (wt.%) | 0.097879 | 0.086255 | 0.1095 | 0.091895 | 0.10386
As, Arsenic (ppm) | 296.0966 | 285.2686 | 306.9245 | 287.9526 | 304.2406
Ba, Barium (ppm) | 33.7921 | 29.1252 | 38.4589 | 31.2321 | 36.3521
Be, Beryllium (ppm) | < 0.5 | IND | IND | IND | IND
Bi, Bismuth (ppm) | 113.2697 | 109.3166 | 117.2227 | 110.1643 | 116.375
Ca, Calcium (wt.%) | 0.32944 | 0.31637 | 0.34252 | 0.32217 | 0.33671
Cd, Cadmium (ppm) | 155.9678 | 149.3449 | 162.5906 | 151.4685 | 160.467
Ce, Cerium (ppm) | 15.7701 | 13.1808 | 18.3594 | 15.1909 | 16.3493
Co, Cobalt (ppm) | 18.3884 | 17.2542 | 19.5225 | 17.6841 | 19.0926
Cr, Chromium (ppm) | 111.6705 | 104.1347 | 119.2063 | 107.6042 | 115.7369
Cs, Caesium (ppm) | 0.14973 | 0.10843 | 0.19103 | IND | IND
Cu, Copper (wt.%) | 0.34712 | 0.33549 | 0.35874 | 0.34001 | 0.35422
Fe, Iron (wt.%) | 2.8834 | 2.7929 | 2.9739 | 2.8331 | 2.9337
Ga, Gallium (ppm) | 1.0046 | 0.86876 | 1.1405 | 0.96665 | 1.0426
In, Indium (ppm) | 1.6342 | 1.4457 | 1.8227 | 1.5463 | 1.7221
La, Lanthanum (ppm) | 9.8864 | 8.4644 | 11.3083 | 9.4576 | 10.3152
Li, Lithium (ppm) | 1.4072 | 1.0998 | 1.7146 | 1.324 | 1.4903
Mg, Magnesium (wt.%) | 0.15221 | 0.14335 | 0.16107 | 0.14903 | 0.1554
Mn, Manganese (wt.%) | 0.136 | 0.13066 | 0.14134 | 0.13363 | 0.13837
Mo, Molybdenum (ppm) | 9.9511 | 8.9022 | 11.0001 | 9.5524 | 10.3499
Ni, Nickel (ppm) | 15.8123 | 14.1997 | 17.4249 | 15.1472 | 16.4773
P, Phosphorus (wt.%) | 0.013295 | 0.011973 | 0.014616 | 0.012919 | 0.01367
Pb, Lead (wt.%) | 70.0908 | 64.4345 | 75.7471 | 69.0552 | 71.1263
S, Sulphur (wt.%) | 14.2259 | 13.8174 | 14.6344 | 13.9808 | 14.471
Sb, Antimony (wt.%) | 0.26592 | 0.25295 | 0.27889 | 0.25937 | 0.27248
Sc, Scandium (ppm) | < 1 | IND | IND | IND | IND
Sn, Tin (ppm) | 12.5212 | 10.1628 | 14.8797 | 11.8419 | 13.2006
Ta, Tantalum (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 0.41286 | 0.34625 | 0.47947 | 0.39169 | 0.43403
Tl, Thallium (ppm) | 1.9218 | 1.7539 | 2.0897 | 1.8254 | 2.0182
U, Uranium (ppm) | 1.0488 | 0.93348 | 1.1641 | 1.0114 | 1.0862
V, Vanadium (ppm) | 5.025 | 4.4836 | 5.5663 | 4.6488 | 5.4011
Zn, Zinc (wt.%) | 2.7934 | 2.7159 | 2.871 | 2.7526 | 2.8343
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 0.099907 | 0.084763 | 0.11505 | 0.092109 | 0.1077
As, Arsenic (ppm) | 291.6764 | 272.7355 | 310.6173 | 278.2198 | 305.133
Ba, Barium (ppm) | 38.6037 | 32.9503 | 44.2571 | 35.3241 | 41.8833
Bi, Bismuth (ppm) | 111.3597 | 104.7233 | 117.9961 | 107.6977 | 115.0217
Ca, Calcium (wt.%) | 0.33709 | 0.3001 | 0.37408 | 0.31444 | 0.35975
Cd, Cadmium (ppm) | 154.345 | 144.1748 | 164.5152 | 148.8603 | 159.8297
Co, Cobalt (ppm) | 18.0871 | 16.0803 | 20.0939 | IND | IND
Cr, Chromium (ppm) | 117.446 | 93.7261 | 141.166 | 107.9454 | 126.9467
Cu, Copper (wt.%) | 0.34822 | 0.3355 | 0.36093 | 0.34204 | 0.35439
Er, Erbium (ppm) | 0.62335 | 0.48031 | 0.76639 | IND | IND
Fe, Iron (wt.%) | 2.9494 | 2.8419 | 3.0569 | 2.8922 | 3.0067
Ga, Gallium (ppm) | 1.0156 | 0.89153 | 1.1397 | IND | IND
In, Indium (ppm) | 1.641 | 1.4758 | 1.8063 | IND | IND
Mg, Magnesium (wt.%) | 0.15585 | 0.14568 | 0.16602 | 0.14948 | 0.16223
Mn, Manganese (wt.%) | 0.13747 | 0.13186 | 0.14308 | 0.13427 | 0.14068
Mo, Molybdenum (ppm) | 11.01 | 9.4085 | 12.6115 | IND | IND
S, Sulphur (wt.%) | 14.6942 | 14.1645 | 15.224 | 14.4894 | 14.8991
Sb, Antimony (wt.%) | 0.27979 | 0.25847 | 0.3011 | 0.27034 | 0.28923
Si, Silicon (wt.%) | 1.4824 | 1.4058 | 1.559 | 1.447 | 1.5178
Sm, Samarium (ppm) | 1.4213 | 1.1379 | 1.7047 | IND | IND
Sr, Strontium (ppm) | < 10 | IND | IND | IND | IND
Tl, Thallium (ppm) | 2.0221 | 1.8393 | 2.2049 | IND | IND
U, Uranium (ppm) | 1.2606 | 1.0601 | 1.4611 | IND | IND
Y, Yttrium (ppm) | 7.7886 | 6.728 | 8.8492 | 7.4415 | 8.1356
Yb, Ytterbium (ppm) | 0.50325 | 0.3622 | 0.64431 | IND | IND
Zn, Zinc (wt.%) | 2.8594 | 2.7574 | 2.9613 | 2.804 | 2.9148
Infrared Combustion
C, Carbon (wt.%) | 0.38655 | 0.36995 | 0.40314 | 0.37504 | 0.39805
S, Sulphur (wt.%) | 14.0336 | 13.6611 | 14.406 | 13.9013 | 14.1659
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 353c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Umpire Labs ('as received' sample basis)
Thermogravimetry
H2O- | wt.% | 0.13037
Geoanalytical Labs ('as received' sample basis)
4-Acid Digestion
Dy | ppm | 0.75937 | Lu | ppm | 0.044311 | Sr | ppm | 2.1489
Er | ppm | 0.35379 | Na | wt.% | 0.014014 | Tb | ppm | 0.15753
Eu | ppm | 1.7253 | Nb | ppm | 0.25738 | Te | ppm | 0.69187
Gd | ppm | 1.0862 | Nd | ppm | 6.4434 | Ti | wt.% | 0.0098133
Ge | ppm | 0.36292 | Pr | ppm | 1.7589 | Tm | ppm | 0.045447
Hf | ppm | 0.13333 | Rb | ppm | 1.2705 | W | ppm | 1.9422
Hg | ppm | 13.3083 | Re | ppm | < 0.05 | Y | ppm | 5.8109
Ho | ppm | 0.11532 | Se | ppm | 5.1579 | Yb | ppm | 0.2838
K | wt.% | 0.020016 | Sm | ppm | 1.2514 | Zr | ppm | 2.5283
Peroxide Fusion ICP
Ag | wt.% | 0.27483 | La | ppm | 9.9611 | Se | ppm | < 20
B | ppm | < 50 | Li | ppm | 5.0372 | Sn | ppm | 15.2889
Be | ppm | < 1 | Lu | ppm | 0.063642 | Ta | ppm | 0.21111
Ce | ppm | 16.5183 | Nb | ppm | < 5 | Tb | ppm | 0.18827
Cs | ppm | 0.20519 | Nd | ppm | 7.0917 | Te | ppm | 4.9833
Dy | ppm | 1.325 | Ni | ppm | 15.7083 | Th | ppm | 0.41403
Eu | ppm | 1.8999 | P | wt.% | 0.020587 | Ti | wt.% | 0.0082587
Gd | ppm | 1.5367 | Pb | wt.% | 71.3327 | Tm | ppm | 0.076133
Ge | ppm | < 1 | Pr | ppm | 1.8915 | V | ppm | < 10
Hf | ppm | 0.1 | Rb | ppm | 1.3509 | W | ppm | 4.9937
Ho | ppm | 0.22648 | Re | ppm | < 0.1 | Zr | ppm | 4.8333
K | wt.% | 0.04201 | Sc | ppm | < 5
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 353c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Umpire Labs (dry sample basis)
Classical Wet Chemistry
Pb, wt.% | 72.2872 | 0.12527 | 72.0367 | 72.5378 | 71.9114 | 72.663 | 0.1733% | 0.3466% | 0.5199% | 68.6729 | 75.9016
Pb Fire Assay
Ag, wt.% | 0.26856 | 0.0021461 | 0.26427 | 0.27286 | 0.26212 | 0.275 | 0.7991% | 1.598% | 2.397% | 0.25513 | 0.28199
Geoanalytical Labs ('as received' sample basis)
4-Acid Digestion
Ag, wt.% | 0.26836 | 0.010208 | 0.24794 | 0.28878 | 0.23774 | 0.29899 | 3.804% | 7.608% | 11.41% | 0.25494 | 0.28178
Al, wt.% | 0.097879 | 0.007881 | 0.082117 | 0.11364 | 0.074236 | 0.12152 | 8.052% | 16.1% | 24.16% | 0.092985 | 0.10277
As, ppm | 296.0966 | 6.5463 | 283.0039 | 309.1892 | 276.4576 | 315.7355 | 2.211% | 4.422% | 6.633% | 281.2917 | 310.9014
Ba, ppm | 33.7921 | 5.8236 | 22.1448 | 45.4393 | 16.3212 | 51.263 | 17.23% | 34.47% | 51.7% | 32.1025 | 35.4817
Be, ppm | < 0.5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Bi, ppm | 113.2697 | 4.1902 | 104.8892 | 121.6501 | 100.6989 | 125.8404 | 3.699% | 7.399% | 11.1% | 107.6062 | 118.9331
Ca, wt.% | 0.32944 | 0.010484 | 0.30848 | 0.35041 | 0.29799 | 0.36089 | 3.182% | 6.364% | 9.547% | 0.31297 | 0.34591
Cd, ppm | 155.9678 | 9.5406 | 136.8865 | 175.049 | 127.3459 | 184.5896 | 6.117% | 12.23% | 18.35% | 148.1694 | 163.7661
Ce, ppm | 15.7701 | 2.1749 | 11.4202 | 20.12 | 9.2453 | 22.295 | 13.79% | 27.58% | 41.37% | 14.9816 | 16.5586
Co, ppm | 18.3884 | 0.8492 | 16.69 | 20.0868 | 15.8408 | 20.936 | 4.618% | 9.236% | 13.85% | 17.4689 | 19.3078
Cr, ppm | 111.6705 | 10.7098 | 90.2508 | 133.0902 | 79.541 | 143.8 | 9.591% | 19.18% | 28.77% | 106.087 | 117.254
Cs, ppm | 0.14973 | 0.021566 | 0.10659 | 0.19286 | 0.085029 | 0.21442 | 14.4% | 28.81% | 43.21% | 0.14224 | 0.15721
Cu, wt.% | 0.34712 | 0.014974 | 0.31717 | 0.37706 | 0.3022 | 0.39204 | 4.314% | 8.627% | 12.94% | 0.32976 | 0.36447
Fe, wt.% | 2.8834 | 0.1148 | 2.6538 | 3.113 | 2.539 | 3.2278 | 3.982% | 7.963% | 11.94% | 2.7392 | 3.0276
Ga, ppm | 1.0046 | 0.16909 | 0.66643 | 1.3428 | 0.49734 | 1.5119 | 16.83% | 33.66% | 50.49% | 0.95437 | 1.0548
In, ppm | 1.6342 | 0.11113 | 1.4119 | 1.8565 | 1.3008 | 1.9676 | 6.8% | 13.6% | 20.4% | 1.5525 | 1.7159
La, ppm | 9.8864 | 1.1614 | 7.5635 | 12.2092 | 6.4021 | 13.3707 | 11.75% | 23.5% | 35.24% | 9.392 | 10.3807
Li, ppm | 1.4072 | 0.16908 | 1.069 | 1.7453 | 0.89993 | 1.9144 | 12.02% | 24.03% | 36.05% | 1.3368 | 1.4775
Mg, wt.% | 0.15221 | 0.0075107 | 0.13719 | 0.16723 | 0.12968 | 0.17474 | 4.934% | 9.869% | 14.8% | 0.1446 | 0.15982
Mn, wt.% | 0.136 | 0.0068734 | 0.12225 | 0.14975 | 0.11538 | 0.15662 | 5.054% | 10.11% | 15.16% | 0.1292 | 0.1428
Mo, ppm | 9.9511 | 1.2644 | 7.4224 | 12.4799 | 6.158 | 13.7443 | 12.71% | 25.41% | 38.12% | 9.4536 | 10.4487
Ni, ppm | 15.8123 | 1.6203 | 12.5716 | 19.0529 | 10.9513 | 20.6732 | 10.25% | 20.49% | 30.74% | 15.0216 | 16.6029
P, wt.% | 0.013295 | 0.0015736 | 0.010148 | 0.016442 | 0.0085739 | 0.018016 | 11.84% | 23.67% | 35.51% | 0.01263 | 0.01396
Pb, wt.% | 70.0908 | 5.0839 | 59.923 | 80.2586 | 54.8391 | 85.3425 | 7.253% | 14.51% | 21.76% | 66.5863 | 73.5953
S, wt.% | 14.2259 | 0.40146 | 13.423 | 15.0288 | 13.0215 | 15.4303 | 2.822% | 5.644% | 8.466% | 13.5146 | 14.9372
Sb, wt.% | 0.26592 | 0.017925 | 0.23007 | 0.30177 | 0.21215 | 0.3197 | 6.741% | 13.48% | 20.22% | 0.25263 | 0.27922
Sc, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sn, ppm | 12.5212 | 2.2915 | 7.9382 | 17.1043 | 5.6466 | 19.3958 | 18.3% | 36.6% | 54.9% | 11.8952 | 13.1473
Ta, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 0.41286 | 0.037799 | 0.33726 | 0.48846 | 0.29947 | 0.52626 | 9.155% | 18.31% | 27.47% | 0.39222 | 0.4335
Tl, ppm | 1.9218 | 0.14793 | 1.626 | 2.2177 | 1.478 | 2.3656 | 7.697% | 15.39% | 23.09% | 1.8257 | 2.0179
U, ppm | 1.0488 | 0.10162 | 0.84554 | 1.252 | 0.74392 | 1.3536 | 9.69% | 19.38% | 29.07% | 0.99634 | 1.1012
V, ppm | 5.025 | 0.60538 | 3.8142 | 6.2357 | 3.2088 | 6.8411 | 12.05% | 24.09% | 36.14% | 4.7737 | 5.2762
Zn, wt.% | 2.7934 | 0.10389 | 2.5857 | 3.0012 | 2.4818 | 3.1051 | 3.719% | 7.438% | 11.16% | 2.6538 | 2.9331
Peroxide Fusion ICP
Al, wt.% | 0.099907 | 0.010421 | 0.079065 | 0.12075 | 0.068645 | 0.13117 | 10.43% | 20.86% | 31.29% | 0.094911 | 0.1049
As, ppm | 291.6764 | 17.4103 | 256.8559 | 326.4969 | 239.4456 | 343.9072 | 5.969% | 11.94% | 17.91% | 277.0926 | 306.2602
Ba, ppm | 38.6037 | 4.5402 | 29.5233 | 47.6841 | 24.9831 | 52.2243 | 11.76% | 23.52% | 35.28% | 36.6735 | 40.5339
Bi, ppm | 111.3597 | 5.2548 | 100.8501 | 121.8693 | 95.5953 | 127.1241 | 4.719% | 9.438% | 14.16% | 105.7917 | 116.9277
Ca, wt.% | 0.33709 | 0.028708 | 0.27968 | 0.39451 | 0.25097 | 0.42322 | 8.516% | 17.03% | 25.55% | 0.32024 | 0.35395
Cd, ppm | 154.345 | 8.3251 | 137.6948 | 170.9952 | 129.3697 | 179.3203 | 5.394% | 10.79% | 16.18% | 146.6277 | 162.0622
Co, ppm | 18.0871 | 2.0026 | 14.082 | 22.0923 | 12.0794 | 24.0949 | 11.07% | 22.14% | 33.22% | 17.1828 | 18.9915
Cr, ppm | 117.446 | 23.3074 | 70.8313 | 164.0608 | 47.5239 | 187.3682 | 19.85% | 39.69% | 59.54% | 111.5737 | 123.3183
Cu, wt.% | 0.34822 | 0.011453 | 0.32531 | 0.37112 | 0.31386 | 0.38257 | 3.289% | 6.578% | 9.867% | 0.3308 | 0.36563
Er, ppm | 0.62335 | 0.086997 | 0.44936 | 0.79734 | 0.36236 | 0.88434 | 13.96% | 27.91% | 41.87% | 0.59218 | 0.65452
Fe, wt.% | 2.9494 | 0.077748 | 2.7939 | 3.1049 | 2.7162 | 3.1827 | 2.636% | 5.272% | 7.908% | 2.8019 | 3.0969
Ga, ppm | 1.0156 | 0.051064 | 0.91349 | 1.1177 | 0.86242 | 1.1688 | 5.028% | 10.06% | 15.08% | 0.96484 | 1.0664
In, ppm | 1.641 | 0.099179 | 1.4427 | 1.8394 | 1.3435 | 1.9386 | 6.044% | 12.09% | 18.13% | 1.559 | 1.7231
Mg, wt.% | 0.15585 | 0.0065835 | 0.14269 | 0.16902 | 0.1361 | 0.1756 | 4.224% | 8.448% | 12.67% | 0.14806 | 0.16364
Mn, wt.% | 0.13747 | 0.0043055 | 0.12886 | 0.14608 | 0.12455 | 0.15039 | 3.132% | 6.264% | 9.396% | 0.1306 | 0.14434
Mo, ppm | 11.01 | 1.3357 | 8.3386 | 13.6813 | 7.003 | 15.017 | 12.13% | 24.26% | 36.39% | 10.4595 | 11.5605
S, wt.% | 14.6942 | 0.54509 | 13.6041 | 15.7844 | 13.059 | 16.3295 | 3.71% | 7.419% | 11.13% | 13.9595 | 15.429
Sb, wt.% | 0.27979 | 0.021755 | 0.23628 | 0.3233 | 0.21452 | 0.34505 | 7.776% | 15.55% | 23.33% | 0.2658 | 0.29378
Si, wt.% | 1.4824 | 0.04076 | 1.4009 | 1.5639 | 1.3601 | 1.6047 | 2.75% | 5.499% | 8.249% | 1.4083 | 1.5565
Sm, ppm | 1.4213 | 0.26051 | 0.9003 | 1.9423 | 0.6398 | 2.2028 | 18.33% | 36.66% | 54.99% | 1.3503 | 1.4924
Sr, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tl, ppm | 2.0221 | 0.10257 | 1.8169 | 2.2272 | 1.7144 | 2.3298 | 5.073% | 10.15% | 15.22% | 1.921 | 2.1232
U, ppm | 1.2606 | 0.15194 | 0.95671 | 1.5645 | 0.80477 | 1.7164 | 12.05% | 24.11% | 36.16% | 1.1976 | 1.3236
Y, ppm | 7.7886 | 1.3682 | 5.0523 | 10.5249 | 3.6841 | 11.893 | 17.57% | 35.13% | 52.7% | 7.3991 | 8.178
Yb, ppm | 0.50325 | 0.070415 | 0.36242 | 0.64408 | 0.29201 | 0.7145 | 13.99% | 27.98% | 41.98% | 0.47809 | 0.52842
Zn, wt.% | 2.8594 | 0.096472 | 2.6664 | 3.0523 | 2.57 | 3.1488 | 3.374% | 6.748% | 10.12% | 2.7164 | 3.0023
Infrared Combustion
C, wt.% | 0.38655 | 0.018935 | 0.34868 | 0.42442 | 0.32974 | 0.44335 | 4.899% | 9.797% | 14.7% | 0.36722 | 0.40587
S, wt.% | 14.0336 | 0.46326 | 13.107 | 14.9601 | 12.6438 | 15.4234 | 3.301% | 6.602% | 9.903% | 13.3319 | 14.7353
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 353c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AmmMo*TITR | Ammonium Molybdate with Titration finish
at 105°C*Oven | moisture at 105°C with drying oven finish
at 105°C*TGA | moisture at 105°C by thermogravimetric analysis
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
EDTA*TITR | ethylenediamminetetraacetic acid titration
FA*GRAV | fire assay with gravimetric finish
IRC | infrared combustion
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*MS*LW | sodium peroxide fusion (low weight) with Mass Spectrometry: ICP-MS finish
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
VOL*TITR | volumetric method with titration
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 353c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AH Knight, Spartanburg, SC, USA
AH Knight, St Helens, Merseyside, UK
AH Knight, Tianjin, China
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
ALS Inspection, Prescot, Merseyside, UK
American Assay Laboratories, Sparks, Nevada, USA
Bachelet, Angleur, Liege, Belgium
Inspectorate (BV), Lima, Peru
Inspectorate (BV), Shanghai, Bao Shan District, China
Inspectorate (BV), Witham, Essex, UK
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek LSI, Rotterdam, Zuid-Holland, Netherlands
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India

## Classical
| Analytical results for Pb in OREAS 353c (Certified Value 72.29 wt.%) | [11 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | wt.%
|  |  | BF*XRF | EDTA*TITR | EDTA*TITR | EDTA*TITR | AmmMo*TITR | VOL*TITR | EDTA*TITR | VOL*TITR | VOL*TITR | EDTA*TITR | 2
3
| 1 | 1 | 71.9852 | 72.5716 | 72.2918 | 72.1878 | 72.2795 | 72.21 | 72.5252 | 72.2941 | 72.1215 | 72.2322 | 1
| 1 | 2 | 71.875 | 72.4614 | 72.2818 | 72.2278 | 72.1894 | 72.19 | 72.447 | 72.2239 | 72.1615 | 72.2523 | 20
| 1 | 3 | 71.8249 | 72.4014 | 72.3219 | 72.2478 | 72.2995 | 72.08 | 72.3328 | 72.2339 | 72.1615 | 72.2723 | 16
| 1 | 4 |  |  |  |  | 72.3296 | 72.27 | 72.6123 |  |  |  | 72.2872
| 1 | 5 | [5 empty] | 72.22 | 72.4942 |  |  |  | 7
| 1 | 6 | [5 empty] | 72.27 | 72.3187
| Mean |  | 71.895 | 72.4781 | 72.2985 | 72.2211 | 72.2745 | 72.2067 | 72.455 | 72.2506 | 72.1482 | 72.2523
| Median |  | 71.875 | 72.4614 | 72.2918 | 72.2278 | 72.2895 | 72.215 | 72.4706 | 72.2339 | 72.1615 | 72.2523
| Std Dev. |  | 0.081985 | 0.086338 | 0.02086 | 0.030575 | 0.060344 | 0.070048 | 0.11381 | 0.037977 | 0.023143 | 0.02002
| Rel.Std.Dev. |  | 0.114% | 0.1191% | 0.02885% | 0.04234% | 0.08349% | 0.09701% | 0.1571% | 0.05256% | 0.03208% | 0.02771%
| PDM3 |  | -0.5425% | 0.2641% | 0.01559% | -0.09146% | -0.0176% | -0.1114% | 0.2321% | -0.0506% | -0.1924% | -0.04838%
| Z-Score (Absolute) |  | 5.23 | 3.33 | 0.69 | 0.45 | 0.34 | 0.66 | 2.99 | 0.01 | 1.52 | 0.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fire Assay
| Analytical results for Ag in OREAS 353c (Certified Value 0.269 wt.%) | [12 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | wt.%
|  |  | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | 3
|  |  | 10g | 10g | 10g | 30g | 30g | 10g | 10g | 10g | 10g | 100g | 15g | 3
| 1 | 1 | 0.26863 | 0.26747 | 0.27187 | 0.26427 | 0.26871 | 0.2683 | 0.2726 | 0.26892 | 0.26753 | 0.26746 | 0.26547 | 1
| 1 | 2 | 0.26903 | 0.27218 | 0.27057 | 0.26355 | 0.26661 | 0.2685 | 0.2712 | 0.2704 | 0.26823 | 0.26716 | 0.26587 | 13
| 1 | 3 | 0.26743 | 0.27098 | 0.27167 | 0.26928 | 0.26841 | 0.26809 | 0.2723 | 0.26818 | 0.26843 | 0.26776 | 0.26727 | 16
| 1 | 4 | [5 empty] | 0.2684 | 0.2699 | 0.26933 |  |  |  | 0.26856
| 1 | 5 | [6 empty] | 0.2709 |  |  |  |  | 9
| Mean |  | 0.26836 | 0.27021 | 0.27137 | 0.2657 | 0.26791 | 0.26832 | 0.27138 | 0.26921 | 0.26806 | 0.26746 | 0.2662
| Median |  | 0.26863 | 0.27098 | 0.27167 | 0.26427 | 0.26841 | 0.26835 | 0.2712 | 0.26912 | 0.26823 | 0.26746 | 0.26587
| Std Dev. |  | 0.000834 | 0.0024454 | 0.00070147 | 0.0031237 | 0.0011367 | 0.00017097 | 0.0010945 | 0.00092811 | 0.00047405 | 0.00030063 | 0.00094611
| Rel.Std.Dev. |  | 0.3108% | 0.905% | 0.2585% | 1.176% | 0.4243% | 0.06372% | 0.4033% | 0.3448% | 0.1768% | 0.1124% | 0.3554%
| PDM3 |  | -0.07456% | 0.6139% | 1.045% | -1.066% | -0.2415% | -0.09041% | 1.049% | 0.2404% | -0.1857% | -0.4101% | -0.88%
| Z-Score (Absolute) |  | 0.03 | 1.44 | 2.31 | 1.99 | 0.31 | 0 | 2.32 | 0.67 | 0.19 | 0.65 | 1.61
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 353c (Indicative Value 0.13 wt.%) | [11 empty] | Indicative
H2O- | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | wt.%
|  |  | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*TGA | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | 3
3
| 1 | 1 | 0.16 | 0.14 | 0.21 | 0.08 | 0.11 | 0.05 | 0.13 | 0.31 | 0.21 | 0.09 | 1
| 1 | 2 | [5 empty] | 0.05 | 0.12 |  |  | 0.11 | 1
| 1 | 3 | [5 empty] | 0.01 | 0.13 |  |  |  | 16
| Mean |  | 0.16 | 0.14 | 0.21 | 0.08 | 0.11 | 0.036667 | 0.12667 | 0.31 | 0.21 | 0.1 | 0.13037
| Median |  | 0.16 | 0.14 | 0.21 | 0.08 | 0.11 | 0.05 | 0.13 | 0.31 | 0.21 | 0.1 | 7
| Std Dev. | [6 empty] | 0.023094 | 0.0057735 |  |  | 0.014142
| Rel.Std.Dev. | [6 empty] | 62.98% | 4.558% |  |  | 14.14%
| PDM3 |  | 22.73% | 7.386% | 61.08% | -38.64% | -15.62% | -71.87% | -2.841% | 137.8% | 61.08% | -23.3%
| Z-Score (Absolute) |  | 0.41 | 0.1 | 1.19 | 0.83 | 0.36 | 1.5 | 0.1 | 2.75 | 1.19 | 0.52
Note on C14: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 353c (Certified Value 0.268 wt.%) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 |  | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.2839 | 0.271 | 0.266 | 0.269 | 0.271 | 0.281 | 0.29888 | 0.2667 | 0.25247 | 0.2452 | 0.26427 | 0.2792 | 0.2619 |  | 1
| 1 | 2 | 0.2752 | 0.273 | 0.275 | 0.27 | 0.275 | 0.277 | 0.29874 | 0.2644 | 0.25575 | 0.25042 | 0.26566 | 0.2685 | 0.2666 |  | 13
| 1 | 3 | 0.2861 | 0.274 | 0.266 | 0.273 | 0.275 | 0.279 | 0.29943 | 0.2675 | 0.25173 | 0.25021 | 0.26282 | 0.2955 | 0.2646 |  | 16
| 1 | 4 | 0.2831 | 0.27 | 0.26 | 0.269 | 0.27 | 0.275 | 0.30015 | 0.2525 | 0.26166 | 0.25181 | 0.2643 | 0.2786 | 0.2704 |  | 0.26836
| 1 | 5 | 0.279 | 0.267 | 0.27 | 0.269 | 0.276 | 0.276 | 0.29906 | 0.2643 | 0.26487 | 0.24802 | 0.26025 | 0.2765 | 0.2655 |  | 11
| 1 | 6 | 0.2904 | 0.269 | 0.267 | 0.274 | 0.27 | 0.277 | 0.29987 | 0.2718 | 0.25871 | 0.24384 | 0.25929 | 0.309 | 0.2649
| Mean |  | 0.28295 | 0.27067 | 0.26733 | 0.27067 | 0.27283 | 0.2775 | 0.29935 | 0.26453 | 0.25753 | 0.24825 | 0.26276 | 0.28455 | 0.26565
| Median |  | 0.2835 | 0.2705 | 0.2665 | 0.2695 | 0.273 | 0.277 | 0.29924 | 0.26555 | 0.25723 | 0.24911 | 0.26354 | 0.2789 | 0.2652
| Std Dev. |  | 0.0053294 | 0.002582 | 0.0049666 | 0.0022509 | 0.0027869 | 0.0021679 | 0.00056487 | 0.0064988 | 0.0051911 | 0.0031632 | 0.0025078 | 0.01486 | 0.0028005
| Rel.Std.Dev. |  | 1.884% | 0.9539% | 1.858% | 0.8316% | 1.021% | 0.7812% | 0.1887% | 2.457% | 2.016% | 1.274% | 0.9544% | 5.222% | 1.054%
| PDM3 |  | 5.436% | 0.8589% | -0.3832% | 0.8589% | 1.666% | 3.405% | 11.55% | -1.427% | -4.036% | -7.494% | -2.086% | 6.032% | -1.01%
| Z-Score (Absolute) |  | 1.21 | 0 | 0.33 | 0 | 0.21 | 0.67 | 2.83 | 0.61 | 1.3 | 2.21 | 0.78 | 1.37 | 0.5
| Analytical results for Al in OREAS 353c (Certified Value 0.098 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.1 | 0.087 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10811 | 0.11 | 0.1 | 0.11 | 0.093758 | 0.1 | 0.0958 |  | 1
| 1 | 2 | 0.11 | 0.088 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10203 | 0.21 | 0.1 | 0.1 | 0.098296 | 0.09 | 0.0994 |  | #N/A
| 1 | 3 | 0.11 | 0.084 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10962 | 0.14 | 0.1 | 0.1 | 0.10199 | 0.1 | 0.0943 |  | 16
| 1 | 4 | 0.11 | 0.086 | 0.1 | 0.1 | 0.09 | 0.1 | 0.10951 | 0.13 | 0.1 | 0.1 | 0.099579 | 0.09 | 0.0961 |  | 0.097879
| 1 | 5 | 0.1 | 0.087 | 0.1 | 0.1 | 0.09 | 0.1 | 0.10972 | 0.14 | 0.1 | 0.1 | 0.090783 | 0.1 | 0.0987 |  | 12
| 1 | 6 | 0.11 | 0.087 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10836 | 0.13 | 0.1 | 0.1 | 0.090043 | 0.09 | 0.0958
| Mean |  | 0.10667 | 0.0865 | 0.1 | 0.1 | 0.096667 | 0.1 | 0.10789 | 0.14333 | 0.1 | 0.10167 | 0.095742 | 0.095 | 0.096683
| Median |  | 0.11 | 0.087 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10894 | 0.135 | 0.1 | 0.1 | 0.096027 | 0.095 | 0.09595
| Std Dev. |  | 0.005164 | 0.0013784 | 1.5202e-17 | 1.5202e-17 | 0.005164 | 1.5202e-17 | 0.002952 | 0.034448 | 1.5202e-17 | 0.0040825 | 0.0049265 | 0.0054772 | 0.0019508
| Rel.Std.Dev. |  | 4.841% | 1.594% | 1.52e-14% | 1.52e-14% | 5.342% | 1.52e-14% | 2.736% | 24.03% | 1.52e-14% | 4.016% | 5.146% | 5.766% | 2.018%
| PDM3 |  | 8.978% | -11.63% | 2.167% | 2.167% | -1.238% | 2.167% | 10.23% | 46.44% | 2.167% | 3.87% | -2.183% | -2.941% | -1.221%
| Z-Score (Absolute) |  | 1.15 | 1.18 | NA | NA | 0 | NA | 1.3 | 5.39 | NA | NA | 0.11 | 0.19 | 0
Results from laboratory 3, 5, 7, 10 and 11 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for As in OREAS 353c (Certified Value 296 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 304.71 | 276 | 301 | 303 | 299 | 298 | 281 | 287 | 294 | 292.6194 | 249 | 294.7 | 296 |  | 1
| 1 | 2 | 310.82 | 275 | 314 | 294 | 298 | 301 | 297 | 293 | 292 | 290.8285 | 250 | 294.9 | 295 |  | #N/A
| 1 | 3 | 317.7 | 269 | 307 | 298 | 299 | 300 | 288 | 301 | 297 | 289.4605 | 262 | 286.2 | 308 |  | 16
| 1 | 4 | 314.28 | 269 | 310 | 292 | 291 | 300 | 292 | 304 | 295 | 286.803 | 244 | 294.2 | 301 |  | 296.0966
| 1 | 5 | 308.92 | 266 | 302 | 290 | 292 | 298 | 283 | 289 | 291 | 299.2345 | 262 | 294 | 300 |  | 13
| 1 | 6 | 309.19 | 269 | 309 | 301 | 302 | 295 | 293 | 297 | 299 | 297.9482 | 260 | 285.9 | 292
| Mean |  | 310.9367 | 270.6667 | 307.1667 | 296.3333 | 296.8333 | 298.6667 | 289 | 295.1667 | 294.6667 | 292.8157 | 254.5 | 291.65 | 298.6667
| Median |  | 310.005 | 269 | 308 | 296 | 298.5 | 299 | 290 | 295 | 294.5 | 291.724 | 255 | 294.1 | 298
| Std Dev. |  | 4.5348 | 3.9328 | 4.9565 | 5.164 | 4.3551 | 2.1602 | 6.1644 | 6.7057 | 3.0111 | 4.8773 | 7.791 | 4.351 | 5.6451
| Rel.Std.Dev. |  | 1.458% | 1.453% | 1.614% | 1.743% | 1.467% | 0.7233% | 2.133% | 2.272% | 1.022% | 1.666% | 3.061% | 1.492% | 1.89%
| PDM3 |  | 5.012% | -8.588% | 3.739% | 0.07996% | 0.2488% | 0.868% | -2.397% | -0.3141% | -0.4829% | -1.108% | -14.05% | -1.502% | 0.868%
| Z-Score (Absolute) |  | 3.04 | 4.72 | 2.31 | 0.22 | 0.32 | 0.67 | 1.19 | 0 | 0.1 | 0.45 | 7.83 | 0.68 | 0.67
| Analytical results for Ba in OREAS 353c (Certified Value 33.8 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
1
| 1 | 1 | 31 | 34 | 40 | 40 | 40 | 40 | 37.82 | 23 | 38 | 32 | 27.9609 | 32 | 40 | 47 | 1
| 1 | 2 | 31 | 34 | 40 | 40 | 40 | 40 | 37.61 | 38 | 38 | 28 | 23.817 | 29 | 43 | 63 | #N/A
| 1 | 3 | 32 | 36 | 40 | 40 | 40 | 40 | 37.88 | 36 | 41 | 31 | 27.0205 | 31 | 41 | 53 | 16
| 1 | 4 | 32 | 34 | 40 | 40 | 40 | 40 | 37.47 | 36 | 39 | 34 | 24.6331 | 22 | 41 | 36 | 33.7921
| 1 | 5 | 30 | 35 | 40 | 40 | 40 | 40 | 37.32 | 41 | 39 | 32 | 27.183 | 40 | 42 | 62 | 14
| 1 | 6 | 31 | 36 | 40 | 40 | 40 | 40 | 37.38 | 18 | 40 | 33 | 23.6771 | 30 | 41 | 53
| Mean |  | 31.1667 | 34.8333 | 40 | 40 | 40 | 40 | 37.58 | 32 | 39.1667 | 31.6667 | 25.7153 | 30.6667 | 41.3333 | 52.3333
| Median |  | 31 | 34.5 | 40 | 40 | 40 | 40 | 37.54 | 36 | 39 | 32 | 25.8268 | 30.5 | 41 | 53
| Std Dev. |  | 0.75277 | 0.98319 | 0 | 0 | 0 | 0 | 0.2316 | 9.2304 | 1.169 | 2.0656 | 1.8883 | 5.785 | 1.0328 | 10.0333
| Rel.Std.Dev. |  | 2.415% | 2.823% | 0% | 0% | 0% | 0% | 0.6163% | 28.84% | 2.985% | 6.523% | 7.343% | 18.86% | 2.499% | 19.17%
| PDM3 |  | -7.769% | 3.081% | 18.37% | 18.37% | 18.37% | 18.37% | 11.21% | -5.303% | 15.9% | -6.29% | -23.9% | -9.249% | 22.32% | 54.87%
| Z-Score (Absolute) |  | 0.44 | 0.28 | NA | NA | NA | NA | 0.81 | 0.28 | 1.12 | 0.34 | 1.5 | 0.54 | 1.54 | 3.69
Results from laboratories 3, 5, 6 and 7 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 353c (Certified Value < 0.5 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.12 | 1.13 | 1.14 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.17 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.12254 | < 1 | < 0.5 | 0.11 | 0.14 |  |  |  | 1
| 1 | 2 | 0.18 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.1304 | < 1 | < 0.5 | 0.13 | 0.14 |  |  |  | 15
| 1 | 3 | 0.18 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.1084 | < 1 | < 0.5 | 0.12 | 0.14 |  |  |  | 16
| 1 | 4 | 0.17 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.12659 | < 1 | < 0.5 | 0.11 | 0.14 |  |  |  | < 0.5
| 1 | 5 | 0.18 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.1142 | < 1 | < 0.5 | 0.17 | 0.14 |  |  |  | 15
| 1 | 6 | 0.17 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.10621 | < 1 | < 0.5 | 0.1 | 0.14
| Mean |  | 0.175 | [5 empty] | 0.11806 |  |  | 0.12333 | 0.14
| Median |  | 0.175 | [5 empty] | 0.11837 |  |  | 0.115 | 0.14
| Std Dev. |  | 0.0054772 | [5 empty] | 0.0099355 |  |  | 0.025033 | 0
| Rel.Std.Dev. |  | 3.13% | [5 empty] | 8.416% |  |  | 20.3% | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Bi in OREAS 353c (Certified Value 113 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 122.28 | 113 | 113 | 115 | 115 | 113 | 107.67 | 102 | 116.5 | 117.5 | 114.1603 | 96.4 | 101.65 | 109.2 | 1
| 1 | 2 | 123.78 | 114 | 116 | 115 | 114 | 110 | 109.29 | 106 | 114 | 117.7 | 116.9555 | 97.9 | 93.64 | 107.2 | #N/A
| 1 | 3 | 111.98 | 114 | 114 | 113 | 114 | 111 | 108.17 | 106 | 116.9 | 119.4 | 114.7324 | 105 | 95.01 | 107 | 16
| 1 | 4 | 118.88 | 114 | 112 | 113 | 110 | 112 | 109.05 | 112 | 113.3 | 117.7 | 114.6597 | 96.9 | 102.75 | 109.9 | 113.2697
| 1 | 5 | 121.22 | 114 | 114 | 110 | 113 | 114 | 109.99 | 105 | 114.7 | 118.4 | 117.2872 | 102 | 94.17 | 109.7 | 16
| 1 | 6 | 119.85 | 114 | 112 | 116 | 117 | 110 | 107.81 | 113 | 112.1 | 119.6 | 116.4498 | 103 | 115.07 | 107.4
| Mean |  | 119.665 | 113.8333 | 113.5 | 113.6667 | 113.8333 | 111.6667 | 108.6633 | 107.3333 | 114.5833 | 118.3833 | 115.7075 | 100.2 | 100.3817 | 108.4
| Median |  | 120.535 | 114 | 113.5 | 114 | 114 | 111.5 | 108.61 | 106 | 114.35 | 118.05 | 115.5911 | 99.95 | 98.33 | 108.3
| Std Dev. |  | 4.1449 | 0.40825 | 1.5166 | 2.1602 | 2.3166 | 1.633 | 0.9231 | 4.274 | 1.8552 | 0.9196 | 1.3451 | 3.5983 | 8.1958 | 1.3401
| Rel.Std.Dev. |  | 3.464% | 0.3586% | 1.336% | 1.901% | 2.035% | 1.462% | 0.8495% | 3.982% | 1.619% | 0.7768% | 1.162% | 3.591% | 8.165% | 1.236%
| PDM3 |  | 5.646% | 0.4976% | 0.2034% | 0.3505% | 0.4976% | -1.415% | -4.067% | -5.241% | 1.16% | 4.515% | 2.152% | -11.54% | -11.38% | -4.299%
| Z-Score (Absolute) |  | 1.18 | 0.05 | 0.02 | 0.02 | 0.05 | 0.37 | 0.96 | 1.22 | 0.19 | 0.93 | 0.41 | 2.61 | 2.57 | 1.01
| Analytical results for Ca in OREAS 353c (Certified Value 0.329 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.34 | 0.3 | 0.33 | 0.34 | 0.32 | 0.32 | 0.29441 | 0.33 | 0.34 | 0.33 | 0.32422 | 0.32 | 0.323 | 0.31 | 1
| 1 | 2 | 0.34 | 0.286 | 0.34 | 0.33 | 0.32 | 0.33 | 0.29887 | 0.36 | 0.34 | 0.33 | 0.32298 | 0.32 | 0.3269 | 0.33 | #N/A
| 1 | 3 | 0.34 | 0.291 | 0.33 | 0.34 | 0.32 | 0.33 | 0.29726 | 0.35 | 0.34 | 0.32 | 0.32102 | 0.33 | 0.3193 | 0.32 | 16
| 1 | 4 | 0.34 | 0.293 | 0.34 | 0.32 | 0.31 | 0.33 | 0.29507 | 0.35 | 0.34 | 0.33 | 0.32052 | 0.31 | 0.3234 | 0.31 | 0.32944
| 1 | 5 | 0.34 | 0.3 | 0.33 | 0.33 | 0.32 | 0.33 | 0.29525 | 0.34 | 0.34 | 0.33 | 0.32272 | 0.32 | 0.33 | 0.32 | 17
| 1 | 6 | 0.35 | 0.295 | 0.33 | 0.33 | 0.32 | 0.32 | 0.29467 | 0.35 | 0.34 | 0.33 | 0.32402 | 0.33 | 0.3218 | 0.32
| Mean |  | 0.34167 | 0.29417 | 0.33333 | 0.33167 | 0.31833 | 0.32667 | 0.29592 | 0.34667 | 0.34 | 0.32833 | 0.32258 | 0.32167 | 0.32407 | 0.31833
| Median |  | 0.34 | 0.294 | 0.33 | 0.33 | 0.32 | 0.33 | 0.29516 | 0.35 | 0.34 | 0.33 | 0.32285 | 0.32 | 0.3232 | 0.32
| Std Dev. |  | 0.0040825 | 0.0054191 | 0.005164 | 0.0075277 | 0.0040825 | 0.005164 | 0.0017614 | 0.010328 | 0 | 0.0040825 | 0.0015245 | 0.0075277 | 0.0038124 | 0.0075277
| Rel.Std.Dev. |  | 1.195% | 1.842% | 1.549% | 2.27% | 1.282% | 1.581% | 0.5952% | 2.979% | 0% | 1.243% | 0.4726% | 2.34% | 1.176% | 2.365%
| PDM3 |  | 3.71% | -10.71% | 1.181% | 0.675% | -3.372% | -0.8427% | -10.18% | 5.228% | 3.205% | -0.3368% | -2.083% | -2.36% | -1.632% | -3.372%
| Z-Score (Absolute) |  | 1.56 | 2.99 | 0.76 | 0.6 | 0.67 | 0.12 | 2.82 | 2.04 | 1.4 | 0.28 | 0.27 | 0.35 | 0.12 | 0.67
| Analytical results for Cd in OREAS 353c (Certified Value 156 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 154.58 | 147 | 164 | 164 | 166 | 161.5 | 148.6548 | 164 | 162.1 | 151 | 138.5606 | 142 | 163.98 | 160.6 | 1
| 1 | 2 | 157.76 | 148 | 169.5 | 160.5 | 165 | 163 | 151.1946 | 170 | 165.8 | 152.4 | 139.6282 | 140 | 134.09 | 158.5 | #N/A
| 1 | 3 | 159.51 | 146 | 164.5 | 163.5 | 165 | 164 | 149.2362 | 175 | 160.8 | 152.4 | 138.5393 | 154 | 147.14 | 160.3 | 16
| 1 | 4 | 158.31 | 145 | 167 | 159.5 | 159.5 | 163.5 | 150.6846 | 168 | 162.6 | 155.2 | 139.1976 | 137 | 152.56 | 162.3 | 155.9678
| 1 | 5 | 157.71 | 142 | 163 | 160.5 | 163.5 | 161 | 153.459 | 160 | 159 | 149.2 | 139.8073 | 148 | 149.12 | 165.7 | 18
| 1 | 6 | 157.74 | 146 | 167.5 | 164.5 | 166 | 160 | 151.5382 | 168 | 164.4 | 151.4 | 138.0012 | 150 | 139.39 | 161.2
| Mean |  | 157.6017 | 145.6667 | 165.9167 | 162.0833 | 164.1667 | 162.1667 | 150.7946 | 167.5 | 162.45 | 151.9333 | 138.9557 | 145.1667 | 147.7133 | 161.4333
| Median |  | 157.75 | 146 | 165.75 | 162 | 165 | 162.25 | 150.9396 | 168 | 162.35 | 151.9 | 138.8791 | 145 | 148.13 | 160.9
| Std Dev. |  | 1.6328 | 2.0656 | 2.4782 | 2.1545 | 2.4631 | 1.5706 | 1.7223 | 5.1284 | 2.4411 | 1.9866 | 0.70379 | 6.5243 | 10.4429 | 2.4328
| Rel.Std.Dev. |  | 1.036% | 1.418% | 1.494% | 1.329% | 1.5% | 0.9685% | 1.142% | 3.062% | 1.503% | 1.308% | 0.5065% | 4.494% | 7.07% | 1.507%
| PDM3 |  | 1.048% | -6.605% | 6.379% | 3.921% | 5.257% | 3.974% | -3.317% | 7.394% | 4.156% | -2.587% | -10.91% | -6.925% | -5.292% | 3.504%
| Z-Score (Absolute) |  | 0.18 | 1.34 | 0.62 | 0.25 | 0.45 | 0.26 | 0.84 | 0.77 | 0.28 | 0.73 | 1.98 | 1.38 | 1.14 | 0.18
| Analytical results for Ce in OREAS 353c (Certified Value 15.8 ppm) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 1
1
| 1 | 1 | 15.5 | 14.0669 | 23.9 | 18.7 | 19.1 | 14.6 | 12.9 | [7 empty] | 1
| 1 | 2 | 15.16 | 14.1555 | 25.1 | 17.2 | 18.9 | 14.5 | 13.1 | [7 empty] | 16
| 1 | 3 | 15.83 | 14.3652 | 24.9 | 18.2 | 19.4 | 15.3 | 12.5 | [7 empty] | 16
| 1 | 4 | 15.77 | 14.2175 | 24.5 | 17.3 | 19.1 | 14 | 12.7 | [7 empty] | 15.7701
| 1 | 5 | 15.88 | 14.5966 | 23.8 | 17.7 | 19.3 | 15 | 13.3 | [7 empty] | 19
| 1 | 6 | 16.21 | 14.2732 | 24.5 | 17.8 | 19 | 15.1 | 13
| Mean |  | 15.725 | 14.2791 | 24.45 | 17.8167 | 19.1333 | 14.75 | 12.9167
| Median |  | 15.8 | 14.2453 | 24.5 | 17.75 | 19.1 | 14.8 | 12.95
| Std Dev. |  | 0.35837 | 0.18567 | 0.52058 | 0.56362 | 0.18619 | 0.47645 | 0.28577
| Rel.Std.Dev. |  | 2.279% | 1.3% | 2.129% | 3.163% | 0.9731% | 3.23% | 2.212%
| PDM3 |  | -0.2862% | -9.455% | 55.04% | 12.98% | 21.33% | -6.469% | -18.09%
| Z-Score (Absolute) |  | 0 | 0.47 | 2.81 | 0.67 | 1.1 | 0.31 | 0.91
| Analytical results for Co in OREAS 353c (Certified Value 18.4 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 17.18 | 19 | 18 | 19 | 19 | 18 | 17.99 | 17.7 | 19 | 20 | 18.7576 | 16.6 | 20.8 | 15 | 1
| 1 | 2 | 17.9 | 18 | 18 | 17 | 18 | 18 | 18.33 | 19.2 | 19 | 20 | 18.9868 | 16.3 | 21 | 20 | #N/A
| 1 | 3 | 17.92 | 19 | 19 | 19 | 19 | 18 | 18.54 | 18.8 | 19 | 20 | 19.1994 | 17.6 | 20.5 | 18 | 16
| 1 | 4 | 18.19 | 19 | 18 | 18 | 18 | 18 | 18.21 | 18.6 | 19 | 21 | 18.6806 | 16.1 | 20.8 | 14 | 18.3884
| 1 | 5 | 17.93 | 18 | 18 | 18 | 18 | 18 | 18.08 | 17.6 | 20 | 20 | 19.1892 | 17.2 | 21.2 | 21 | 20
| 1 | 6 | 18.2 | 19 | 19 | 19 | 19 | 17 | 18.67 | 18.4 | 20 | 21 | 19.0081 | 17.4 | 20.6 | 18
| Mean |  | 17.8867 | 18.6667 | 18.3333 | 18.3333 | 18.5 | 17.8333 | 18.3033 | 18.3833 | 19.3333 | 20.3333 | 18.9703 | 16.8667 | 20.8167 | 17.6667
| Median |  | 17.925 | 19 | 18 | 18.5 | 18.5 | 18 | 18.27 | 18.5 | 19 | 20 | 18.9975 | 16.9 | 20.8 | 18
| Std Dev. |  | 0.37222 | 0.5164 | 0.5164 | 0.8165 | 0.54772 | 0.40825 | 0.26379 | 0.62743 | 0.5164 | 0.5164 | 0.21507 | 0.6186 | 0.25626 | 2.7325
| Rel.Std.Dev. |  | 2.081% | 2.766% | 2.817% | 4.454% | 2.961% | 2.289% | 1.441% | 3.413% | 2.671% | 2.54% | 1.134% | 3.668% | 1.231% | 15.47%
| PDM3 |  | -2.728% | 1.514% | -0.2992% | -0.2992% | 0.6071% | -3.018% | -0.4624% | -0.02733% | 5.139% | 10.58% | 3.165% | -8.275% | 13.21% | -3.925%
| Z-Score (Absolute) |  | 0.64 | 0.42 | 0.03 | 0.03 | 0.19 | 0.71 | 0.07 | 0.03 | 1.32 | 2.67 | 0.83 | 2.02 | 3.33 | 0.94
| Analytical results for Cr in OREAS 353c (Certified Value 112 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 110 | 125 | 126 | 118 | 112 | 115 | 99.38 | 99 | 107 | 114 | 126.7274 | 77 | 98.6 | 89 | 1
| 1 | 2 | 112 | 121 | 128 | 116 | 114 | 116 | 99.46 | 122 | 113 | 106 | 124.8946 | 70.8 | 99.1 | 93 | #N/A
| 1 | 3 | 114 | 120 | 125 | 118 | 113 | 117 | 99.58 | 112 | 108 | 109 | 127.2722 | 85.5 | 98.9 | 98 | 16
| 1 | 4 | 117 | 123 | 128 | 115 | 111 | 115 | 100.04 | 93 | 111 | 106 | 128.1486 | 69.7 | 98.6 | 92 | 111.6705
| 1 | 5 | 112 | 123 | 125 | 116 | 114 | 113 | 97.53 | 109 | 109 | 109 | 124.5012 | 82.2 | 102.2 | 88 | 21
| 1 | 6 | 114 | 120 | 129 | 121 | 114 | 116 | 101.42 | 108 | 110 | 117 | 127.5464 | 90 | 98.4 | 90
| Mean |  | 113.1667 | 122 | 126.8333 | 117.3333 | 113 | 115.3333 | 99.5683 | 107.1667 | 109.6667 | 110.1667 | 126.515 | 79.2 | 99.3 | 91.6667
| Median |  | 113 | 122 | 127 | 117 | 113.5 | 115.5 | 99.52 | 108.5 | 109.5 | 109 | 126.9998 | 79.6 | 98.75 | 91
| Std Dev. |  | 2.4014 | 2 | 1.7224 | 2.1602 | 1.2649 | 1.3663 | 1.2533 | 10.1472 | 2.1602 | 4.446 | 1.4854 | 8.1361 | 1.4422 | 3.6148
| Rel.Std.Dev. |  | 2.122% | 1.639% | 1.358% | 1.841% | 1.119% | 1.185% | 1.259% | 9.469% | 1.97% | 4.036% | 1.174% | 10.27% | 1.452% | 3.943%
| PDM3 |  | 1.34% | 9.25% | 13.58% | 5.071% | 1.191% | 3.28% | -10.84% | -4.033% | -1.794% | -1.347% | 13.29% | -29.08% | -11.08% | -17.91%
| Z-Score (Absolute) |  | 0.13 | 0.87 | 1.27 | 0.48 | 0.12 | 0.31 | 1 | 0.37 | 0.16 | 0.12 | 1.25 | 2.7 | 1.02 | 1.66
| Analytical results for Cs in OREAS 353c (Certified Value 0.15 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.12 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | 0.15 | 0.2 | 0.11088 | < 1 | 0.15738 | 0.18 | < 0.2 | [7 empty] | 1
| 1 | 2 | 0.16 | 0.2 | 0.11781 | < 1 | 0.154 | 0.17 | < 0.2 | [7 empty] | 17
| 1 | 3 | 0.15 | 0.2 | 0.11661 | < 1 | 0.15328 | 0.17 | < 0.2 | [7 empty] | 16
| 1 | 4 | 0.16 | 0.2 | 0.11064 | < 1 | 0.15643 | 0.16 | < 0.2 | [7 empty] | 0.14973
| 1 | 5 | 0.15 | 0.2 | 0.1186 | < 1 | 0.15581 | 0.18 | 0.2 | [7 empty] | 22
| 1 | 6 | 0.15 | 0.2 | 0.12751 | < 1 | 0.15448 | 0.18 | 0.2
| Mean |  | 0.15333 | 0.2 | 0.11701 |  | 0.15523 | 0.17333 | 0.2
| Median |  | 0.15 | 0.2 | 0.11721 |  | 0.15514 | 0.175 | 0.2
| Std Dev. |  | 0.005164 | 3.0405e-17 | 0.0061937 |  | 0.0015681 | 0.008165 | 0
| Rel.Std.Dev. |  | 3.368% | 1.52e-14% | 5.293% |  | 1.01% | 4.711% | 0%
| PDM3 |  | 2.409% | 33.58% | -21.85% |  | 3.676% | 15.77% | 33.58%
| Z-Score (Absolute) |  | 0.07 | NA | 1.42 | 12.84 | 0 | 0.67 | NA
Results from laboratories 2 and 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 353c (Certified Value 0.347 wt.%) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.36738 | 0.347 | 0.339 | 0.337 | 0.338 | 0.337 | 0.35851 | 0.3699 | 0.3458 | 0.3489 | 0.34861 | 0.318 | 0.36637 | 0.3515 | 1
| 1 | 2 | 0.37773 | 0.344 | 0.35 | 0.329 | 0.333 | 0.342 | 0.35803 | 0.3633 | 0.3418 | 0.3448 | 0.35401 | 0.314 | 0.37284 | 0.352 | #N/A
| 1 | 3 | 0.37666 | 0.345 | 0.341 | 0.334 | 0.335 | 0.342 | 0.35946 | 0.3647 | 0.3453 | 0.3381 | 0.35264 | 0.331 | 0.36097 | 0.347 | 16
| 1 | 4 | 0.37038 | 0.341 | 0.344 | 0.326 | 0.324 | 0.343 | 0.35413 | 0.3654 | 0.3463 | 0.3431 | 0.35261 | 0.305 | 0.36528 | 0.3577 | 0.34712
| 1 | 5 | 0.37037 | 0.342 | 0.338 | 0.327 | 0.33 | 0.337 | 0.35174 | 0.3665 | 0.341 | 0.3459 | 0.34854 | 0.321 | 0.37239 | 0.343 | 23
| 1 | 6 | 0.3735 | 0.345 | 0.346 | 0.339 | 0.336 | 0.335 | 0.33432 | 0.3691 | 0.3425 | 0.3425 | 0.35617 | 0.328 | 0.36247 | 0.3536
| Mean |  | 0.37267 | 0.344 | 0.343 | 0.332 | 0.33267 | 0.33933 | 0.3527 | 0.36648 | 0.34378 | 0.34388 | 0.35209 | 0.3195 | 0.36672 | 0.3508
| Median |  | 0.37194 | 0.3445 | 0.3425 | 0.3315 | 0.334 | 0.3395 | 0.35608 | 0.36595 | 0.3439 | 0.34395 | 0.35262 | 0.3195 | 0.36583 | 0.35175
| Std Dev. |  | 0.0040166 | 0.0021909 | 0.0045607 | 0.0054406 | 0.0050465 | 0.0033862 | 0.0094732 | 0.0025694 | 0.0022816 | 0.0036345 | 0.0030199 | 0.0094816 | 0.0049583 | 0.0051509
| Rel.Std.Dev. |  | 1.078% | 0.6369% | 1.33% | 1.639% | 1.517% | 0.9979% | 2.686% | 0.7011% | 0.6637% | 1.057% | 0.8577% | 2.968% | 1.352% | 1.468%
| PDM3 |  | 7.361% | -0.8979% | -1.186% | -4.355% | -4.163% | -2.242% | 1.608% | 5.579% | -0.9603% | -0.9315% | 1.434% | -7.956% | 5.647% | 1.061%
| Z-Score (Absolute) |  | 2.29 | 0 | 0.08 | 0.95 | 0.9 | 0.37 | 0.7 | 1.8 | 0.01 | 0 | 0.65 | 1.95 | 1.82 | 0.55
| Analytical results for Dy in OREAS 353c (Indicative Value 0.76 ppm) | [15 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.47 | 0.66483 | 1 | 1.2 | [10 empty] | 1
| 1 | 2 | 0.49 | 0.64121 | 0.9 | 0.9 | [10 empty] | 3
| 1 | 3 | 0.48 | 0.67643 | 1 | 1 | [10 empty] | 16
| 1 | 4 | 0.47 | 0.63259 | 0.8 | 1 | [10 empty] | 0.75937
| 1 | 5 | 0.5 | 0.68007 | 0.9 | 1 | [10 empty] | 9
| 1 | 6 | 0.48 | 0.63986 | 0.8 | 0.9
| Mean |  | 0.48167 | 0.65583 | 0.9 | 1
| Median |  | 0.48 | 0.65302 | 0.9 | 1
| Std Dev. |  | 0.01169 | 0.020503 | 0.089443 | 0.10954
| Rel.Std.Dev. |  | 2.427% | 3.126% | 9.938% | 10.95%
| PDM3 |  | -36.57% | -13.64% | 18.52% | 31.69%
| Z-Score (Absolute) |  | 1.16 | 0.48 | 0.48 | 0.87
| Analytical results for Er in OREAS 353c (Indicative Value 0.35 ppm) | [15 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.2 | 0.34017 | 0.4 | 0.5 | [10 empty] | 1
| 1 | 2 | 0.19 | 0.34929 | 0.4 | 0.4 | [10 empty] | 4
| 1 | 3 | 0.2 | 0.35143 | 0.5 | 0.4 | [10 empty] | 16
| 1 | 4 | 0.2 | 0.36554 | 0.6 | 0.4 | [10 empty] | 0.35379
| 1 | 5 | 0.19 | 0.34434 | 0.4 | 0.4 | [10 empty] | 10
| 1 | 6 | 0.21 | 0.30232 | 0.4 | 0.4
| Mean |  | 0.19833 | 0.34218 | 0.45 | 0.41667
| Median |  | 0.2 | 0.34682 | 0.4 | 0.4
| Std Dev. |  | 0.0075277 | 0.021349 | 0.083666 | 0.040825
| Rel.Std.Dev. |  | 3.795% | 6.239% | 18.59% | 9.798%
| PDM3 |  | -43.94% | -3.281% | 27.19% | 17.77%
| Z-Score (Absolute) |  | 2.27 | 0.47 | 0.88 | 0.47
| Analytical results for Eu in OREAS 353c (Indicative Value 1.73 ppm) | [15 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 1.16 | 1.4233 | 2 | 2.1 | [10 empty] | 1
| 1 | 2 | 1.18 | 1.4151 | 2.1 | 2.2 | [10 empty] | 5
| 1 | 3 | 1.15 | 1.3851 | 2.3 | 2.2 | [10 empty] | 16
| 1 | 4 | 1.14 | 1.4 | 2 | 2.3 | [10 empty] | 1.7253
| 1 | 5 | 1.1 | 1.3904 | 2.2 | 2.3 | [10 empty] | 11
| 1 | 6 | 1.15 | 1.4127 | 2.1 | 2.3
| Mean |  | 1.1467 | 1.4044 | 2.1167 | 2.2333
| Median |  | 1.15 | 1.4063 | 2.1 | 2.25
| Std Dev. |  | 0.026583 | 0.015013 | 0.1169 | 0.08165
| Rel.Std.Dev. |  | 2.318% | 1.069% | 5.523% | 3.656%
| PDM3 |  | -33.54% | -18.6% | 22.69% | 29.45%
| Z-Score (Absolute) |  | 1 | 0.58 | 0.58 | 0.77
| Analytical results for Fe in OREAS 353c (Certified Value 2.88 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.93 | 2.86 | 2.85 | 2.85 | 2.87 | 2.82 | 2.7674 | 3.04 | 2.97 | 2.94 | 2.9954 | 2.7 | 2.7217 | 3.05 | 1
| 1 | 2 | 2.95 | 2.89 | 2.92 | 2.78 | 2.83 | 2.87 | 2.7694 | 3.21 | 2.97 | 2.91 | 3.0175 | 2.73 | 2.7765 | 2.96 | #N/A
| 1 | 3 | 2.95 | 2.88 | 2.88 | 2.85 | 2.85 | 2.84 | 2.7672 | 3.15 | 2.97 | 2.84 | 3.0246 | 2.85 | 2.704 | 2.93 | 16
| 1 | 4 | 2.93 | 2.82 | 2.89 | 2.78 | 2.78 | 2.87 | 2.7361 | 3.13 | 2.99 | 2.9 | 2.9883 | 2.61 | 2.727 | 2.94 | 2.8834
| 1 | 5 | 2.93 | 2.89 | 2.84 | 2.79 | 2.8 | 2.81 | 2.7281 | 3.05 | 2.95 | 2.9 | 3.055 | 2.74 | 2.8049 | 2.92 | 24
| 1 | 6 | 2.93 | 2.84 | 2.89 | 2.86 | 2.85 | 2.77 | 2.7983 | 3.19 | 2.96 | 2.88 | 3.043 | 2.76 | 2.7308 | 2.97
| Mean |  | 2.9367 | 2.8633 | 2.8783 | 2.8183 | 2.83 | 2.83 | 2.7611 | 3.1283 | 2.9683 | 2.895 | 3.0206 | 2.7317 | 2.7441 | 2.9617
| Median |  | 2.93 | 2.87 | 2.885 | 2.82 | 2.84 | 2.83 | 2.7673 | 3.14 | 2.97 | 2.9 | 3.021 | 2.735 | 2.7289 | 2.95
| Std Dev. |  | 0.010328 | 0.028752 | 0.029269 | 0.038687 | 0.034059 | 0.038471 | 0.025461 | 0.070545 | 0.013292 | 0.033317 | 0.026023 | 0.078337 | 0.038277 | 0.047081
| Rel.Std.Dev. |  | 0.3517% | 1.004% | 1.017% | 1.373% | 1.203% | 1.359% | 0.9221% | 2.255% | 0.4478% | 1.151% | 0.8615% | 2.868% | 1.395% | 1.59%
| PDM3 |  | 1.848% | -0.6957% | -0.1755% | -2.256% | -1.852% | -1.852% | -4.242% | 8.495% | 2.946% | 0.4025% | 4.759% | -5.262% | -4.829% | 2.715%
| Z-Score (Absolute) |  | 0.57 | 0.06 | 0.06 | 0.45 | 0.35 | 0.35 | 0.94 | 2.22 | 0.84 | 0.21 | 1.29 | 1.2 | 1.09 | 0.78
| Analytical results for Ga in OREAS 353c (Certified Value 1 ppm) | [15 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 1.11 | 0.9 | < 10 | < 10 | < 10 | < 10 | 0.8182 | 1.1 | 0.7 | 0.8 | 0.84961 | 1.19 | < 0.05 |  | 1
| 1 | 2 | 1.07 | 0.9 | < 10 | < 10 | < 10 | < 10 | 0.81983 | 1.2 | 1 | 0.9 | 0.83965 | 1.21 | < 0.05 |  | #N/A
| 1 | 3 | 1.1 | 0.9 | < 10 | < 10 | < 10 | < 10 | 0.79587 | 1.2 | 0.8 | 1 | 0.85517 | 1.21 | < 0.05 |  | 16
| 1 | 4 | 1.14 | 0.9 | < 10 | < 10 | < 10 | < 10 | 0.81732 | 1.2 | 1 | 1.4 | 0.84537 | 1.17 | < 0.05 |  | 1.0046
| 1 | 5 | 1.09 | 0.9 | < 10 | < 10 | < 10 | < 10 | 0.83406 | 1.1 | 1.2 | 0.9 | 0.8562 | 1.26 | < 0.05 |  | 25
| 1 | 6 | 1.07 | 0.9 | < 10 | < 10 | < 10 | < 10 | 0.81553 | 1.2 | 1.1 | 1.2 | 0.84422 | 1.21 | < 0.05
| Mean |  | 1.0967 | 0.9 |  |  |  |  | 0.8168 | 1.1667 | 0.96667 | 1.0333 | 0.84837 | 1.2083
| Median |  | 1.095 | 0.9 |  |  |  |  | 0.81776 | 1.2 | 1 | 0.95 | 0.84749 | 1.21
| Std Dev. |  | 0.026583 | 0 |  |  |  |  | 0.012238 | 0.05164 | 0.18619 | 0.22509 | 0.0065005 | 0.029944
| Rel.Std.Dev. |  | 2.424% | 0% |  |  |  |  | 1.498% | 4.426% | 19.26% | 21.78% | 0.7662% | 2.478%
| PDM3 |  | 9.164% | -10.41% |  |  |  |  | -18.69% | 16.13% | -3.776% | 2.86% | -15.55% | 20.28%
| Z-Score (Absolute) |  | 0 | 0.53 | 10.6 | 10.6 | 10.6 | 10.6 | 0.76 | 0.19 | 0.35 | 0.17 | 0.67 | 0.3 | 2.91
| Analytical results for Gd in OREAS 353c (Indicative Value 1.09 ppm) | [15 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.72 | 1.1696 | 1.2 | 1.2 | [10 empty] | 1
| 1 | 2 | 0.69 | 1.1568 | 1.3 | 1.3 | [10 empty] | 6
| 1 | 3 | 0.71 | 1.1291 | 1.2 | 1.3 | [10 empty] | 16
| 1 | 4 | 0.7 | 1.1206 | 1.2 | 1.3 | [10 empty] | 1.0862
| 1 | 5 | 0.74 | 1.1529 | 1.2 | 1.2 | [10 empty] | 12
| 1 | 6 | 0.72 | 1.1593 | 1.2 | 1.3
| Mean |  | 0.71333 | 1.148 | 1.2167 | 1.2667
| Median |  | 0.715 | 1.1548 | 1.2 | 1.3
| Std Dev. |  | 0.017512 | 0.019012 | 0.040825 | 0.05164
| Rel.Std.Dev. |  | 2.455% | 1.656% | 3.355% | 4.077%
| PDM3 |  | -34.33% | 5.695% | 12.01% | 16.62%
| Z-Score (Absolute) |  | 5.33 | 0.39 | 0.39 | 0.96
| Analytical results for Ge in OREAS 353c (Indicative Value 0.36 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.12 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.32 | 0.6 | < 0.1 | 0.22 | 0.32 | [9 empty] | 1
| 1 | 2 | 0.34 | 0.6 | < 0.1 | 0.23 | 0.29 | [9 empty] | 7
| 1 | 3 | 0.31 | 0.6 | < 0.1 | 0.26 | 0.29 | [9 empty] | 16
| 1 | 4 | 0.34 | 0.6 | < 0.1 | 0.23 | 0.35 | [9 empty] | 0.36292
| 1 | 5 | 0.32 | 0.6 | < 0.1 | 0.21 | 0.26 | [9 empty] | 13
| 1 | 6 | 0.32 | 0.6 | < 0.1 | 0.24 | 0.26
| Mean |  | 0.325 | 0.6 |  | 0.23167 | 0.295
| Median |  | 0.32 | 0.6 |  | 0.23 | 0.29
| Std Dev. |  | 0.012247 | 0 |  | 0.017224 | 0.035071
| Rel.Std.Dev. |  | 3.768% | 0% |  | 7.435% | 11.89%
| PDM3 |  | -10.45% | 65.33% |  | -36.17% | -18.71%
| Z-Score (Absolute) |  | 0.32 | 3.25 | 2.61 | 0.67 | 0
| Analytical results for Hf in OREAS 353c (Indicative Value 0.13 ppm) | [15 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.12 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 0.29 | < 0.1 | 0.1 | < 0.1 | 0.06 | 0.06 | [8 empty] | 1
| 1 | 2 | 0.29 | < 0.1 | 0.11 | < 0.1 | 0.07 | 0.07 | [8 empty] | 8
| 1 | 3 | 0.3 | < 0.1 | 0.12 | < 0.1 | 0.06 | 0.06 | [8 empty] | 16
| 1 | 4 | 0.3 | < 0.1 | 0.1 | < 0.1 | 0.06 | 0.06 | [8 empty] | 0.13333
| 1 | 5 | 0.29 | < 0.1 | 0.1 | < 0.1 | 0.07 | 0.06 | [8 empty] | 14
| 1 | 6 | 0.3 | < 0.1 | 0.12 | < 0.1 | 0.08 | 0.07
| Mean |  | 0.295 |  | 0.10833 |  | 0.066667 | 0.063333
| Median |  | 0.295 |  | 0.105 |  | 0.065 | 0.06
| Std Dev. |  | 0.0054772 |  | 0.0098319 |  | 0.008165 | 0.005164
| Rel.Std.Dev. |  | 1.857% |  | 9.076% |  | 12.25% | 8.154%
| PDM3 |  | 121.3% |  | -18.75% |  | -50% | -52.5%
| Z-Score (Absolute) |  | 10.34 | 0.67 | 1.95 | 0.67 | 0.07 | 0.07
| Analytical results for Hg in OREAS 353c (Indicative Value 13.3 ppm) | [15 empty] | Indicative
Hg | Round | Replicate | Lab | [13 empty] | 1
| No | No | 1.14 | [13 empty] | ppm
|  |  | 4A*OES/MS | [13 empty] | 1
1
| 1 | 1 | 13.11 | [13 empty] | 1
| 1 | 2 | 13.63 | [13 empty] | 9
| 1 | 3 | 13.03 | [13 empty] | 16
| 1 | 4 | 13.69 | [13 empty] | 13.3083
| 1 | 5 | 13.06 | [13 empty] | 15
| 1 | 6 | 13.33
| Mean |  | 13.3083
| Median |  | 13.22
| Std Dev. |  | 0.2926
| Rel.Std.Dev. |  | 2.199%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 353c (Indicative Value 0.12 ppm) | [15 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.07 | 0.10678 | 0.1 | 0.2 | [10 empty] | 1
| 1 | 2 | 0.07 | 0.12353 | 0.2 | 0.1 | [10 empty] | 10
| 1 | 3 | 0.07 | 0.12165 | 0.1 | 0.2 | [10 empty] | 16
| 1 | 4 | 0.08 | 0.12669 | 0.1 | 0.1 | [10 empty] | 0.11532
| 1 | 5 | 0.08 | 0.11247 | 0.1 | 0.2 | [10 empty] | 16
| 1 | 6 | 0.08 | 0.12657 | 0.1 | 0.1
| Mean |  | 0.075 | 0.11961 | 0.11667 | 0.15
| Median |  | 0.075 | 0.12259 | 0.1 | 0.15
| Std Dev. |  | 0.0054772 | 0.008168 | 0.040825 | 0.054772
| Rel.Std.Dev. |  | 7.303% | 6.829% | 34.99% | 36.51%
| PDM3 |  | -34.96% | 3.724% | 1.167% | 30.07%
| Z-Score (Absolute) |  | 1.75 | 0.06 | 0.06 | 1.29
| Analytical results for In in OREAS 353c (Certified Value 1.63 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  |  | 2
3
| 1 | 1 | 0.93 | 1.6 | 1.2086 | 1.69 | 1.65 | 1.76 | 1.432 | 1.61 | 1.849 | 2 |  |  |  |  | 1
| 1 | 2 | 0.9 | 1.57 | 1.1861 | 1.77 | 1.75 | 1.71 | 1.4546 | 1.64 | 1.465 | 2 |  |  |  |  | #N/A
| 1 | 3 | 0.94 | 1.59 | 1.1291 | 1.79 | 1.64 | 1.67 | 1.4556 | 1.68 | 1.644 | 2 |  |  |  |  | 16
| 1 | 4 | 0.99 | 1.62 | 1.0804 | 1.71 | 1.57 | 1.79 | 1.4369 | 1.55 | 1.721 | 2 |  |  |  |  | 1.6342
| 1 | 5 | 1 | 1.65 | 1.1909 | 1.63 | 1.57 | 1.75 | 1.4605 | 1.69 | 1.623 | 2 |  |  |  |  | 26
| 1 | 6 | 0.99 | 1.58 | 1.0358 | 1.71 | 1.82 | 1.65 | 1.4334 | 1.7 | 1.551 | 2
| Mean |  | 0.95833 | 1.6017 | 1.1385 | 1.7167 | 1.6667 | 1.7217 | 1.4455 | 1.645 | 1.6422 | 2
| Median |  | 0.965 | 1.595 | 1.1576 | 1.71 | 1.645 | 1.73 | 1.4457 | 1.66 | 1.6335 | 2
| Std Dev. |  | 0.040702 | 0.029269 | 0.069216 | 0.057504 | 0.10013 | 0.054559 | 0.012729 | 0.057533 | 0.13347 | 0
| Rel.Std.Dev. |  | 4.247% | 1.827% | 6.08% | 3.35% | 6.008% | 3.169% | 0.8806% | 3.497% | 8.127% | 0%
| PDM3 |  | -41.36% | -1.99% | -30.33% | 5.047% | 1.987% | 5.353% | -11.55% | 0.6614% | 0.488% | 22.38%
| Z-Score (Absolute) |  | 6.19 | 0.37 | 4.56 | 0.67 | 0.22 | 0.72 | 1.78 | 0.03 | 0 | NA
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for K in OREAS 353c (Indicative Value 0.02 wt.%) | [15 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.02 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015588 | 0.03 |  |  | 0.019162 | 0.03 | < 0.03 |  | 1
| 1 | 2 | 0.02 | 0.017 | 0.02 | 0.02 | 0.02 | 0.02 | 0.014738 | 0.04 |  | 0.03 | 0.020619 | 0.02 | < 0.03 |  | 11
| 1 | 3 | 0.02 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016272 | 0.04 |  | 0.03 | 0.019988 | 0.02 | < 0.03 |  | 16
| 1 | 4 | 0.02 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015957 | 0.04 | 0.02 |  | 0.018967 | 0.02 | < 0.03 |  | 0.020016
| 1 | 5 | 0.02 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015613 | 0.03 |  |  | 0.02069 | 0.03 | < 0.03 |  | 17
| 1 | 6 | 0.02 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016565 | 0.03 | 0.02 | 0.02 | 0.019022 | 0.02 | < 0.03
| Mean |  | 0.02 | 0.017833 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015789 | 0.035 | 0.02 | 0.026667 | 0.019741 | 0.023333
| Median |  | 0.02 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015785 | 0.035 | 0.02 | 0.03 | 0.019575 | 0.02
| Std Dev. |  | 0 | 0.00040825 | 0 | 0 | 0 | 0 | 0.00063855 | 0.0054772 | 0 | 0.0057735 | 0.00079796 | 0.005164
| Rel.Std.Dev. |  | 0% | 2.289% | 0% | 0% | 0% | 0% | 4.044% | 15.65% | 0% | 21.65% | 4.042% | 22.13%
| PDM3 |  | -0.0819% | -10.91% | -0.0819% | -0.0819% | -0.0819% | -0.0819% | -21.12% | 74.86% | -0.0819% | 33.22% | -1.374% | 16.57%
| Z-Score (Absolute) |  | 0 | 5.65 | 0 | 0 | 0 | 0 | 10.98 | 39.11 | 0 | 17.38 | 0.67 | 8.69 | 13.04
| Analytical results for La in OREAS 353c (Certified Value 9.89 ppm) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 9.71 | 10 | 10 | 10 | 10 | 8.075 | 13.9 | 10.3 | 11.8 | 9 | 9.9 |  |  |  | 1
| 1 | 2 | 9.4 | 10 | 10 | 10 | 10 | 8.3895 | 14.4 | 9.9 | 11.5 | 8.9 | 10.1 |  |  |  | 13
| 1 | 3 | 9.61 | 10 | 10 | 10 | 10 | 8.2365 | 14.6 | 11 | 12.1 | 9.3 | 9.8 |  |  |  | 16
| 1 | 4 | 10.01 | 10 | 10 | 10 | 10 | 8.19 | 14.1 | 10 | 11.7 | 8.7 | 10 |  |  |  | 9.8864
| 1 | 5 | 9.95 | 10 | 10 | 10 | 10 | 8.3045 | 13.8 | 10.6 | 12.2 | 9.3 | 10.2 |  |  |  | 27
| 1 | 6 | 9.78 | 10 | 10 | 10 | 10 | 8.2535 | 14.2 | 10.5 | 11.9 | 9.3 | 10
| Mean |  | 9.7433 | 10 | 10 | 10 | 10 | 8.2415 | 14.1667 | 10.3833 | 11.8667 | 9.0833 | 10
| Median |  | 9.745 | 10 | 10 | 10 | 10 | 8.245 | 14.15 | 10.4 | 11.85 | 9.15 | 10
| Std Dev. |  | 0.22447 | 0 | 0 | 0 | 0 | 0.10622 | 0.30111 | 0.40702 | 0.2582 | 0.25626 | 0.14142
| Rel.Std.Dev. |  | 2.304% | 0% | 0% | 0% | 0% | 1.289% | 2.125% | 3.92% | 2.176% | 2.821% | 1.414%
| PDM3 |  | -1.447% | 1.149% | 1.149% | 1.149% | 1.149% | -16.64% | 43.3% | 5.027% | 20.03% | -8.123% | 1.149%
| Z-Score (Absolute) |  | 0.19 | NA | NA | NA | NA | 1.29 | 3.07 | 0.28 | 1.37 | 0.67 | 0
Results from laboratories 3, 5, 6 and 7 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 353c (Certified Value 1.41 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
2
| 1 | 1 | 1.49 | 1.6 | < 10 | < 10 | < 10 | < 10 | 1.15 | 2 | 2 | 2 | 1.3813 | 2.2 | < 2 |  | 1
| 1 | 2 | 1.52 | 1.5 | < 10 | < 10 | < 10 | < 10 | 1.1789 | 2 | 1 | 1 | 1.4022 | 2.5 | < 2 |  | 14
| 1 | 3 | 1.55 | 1.5 | < 10 | < 10 | < 10 | < 10 | 1.1647 | 2 | 1 | 2 | 1.4197 | 2.3 | < 2 | 4 | 16
| 1 | 4 | 1.53 | 1.6 | < 10 | < 10 | < 10 | < 10 | 1.117 | 2 | 1 | 2 | 1.3895 | 1.8 | < 2 |  | 1.4072
| 1 | 5 | 1.56 | 1.5 | < 10 | < 10 | < 10 | < 10 | 1.1074 | 2 | 1 | 2 | 1.4299 | 2.1 | < 2 |  | 28
| 1 | 6 | 1.55 | 1.6 | < 10 | < 10 | < 10 | < 10 | 1.1462 | 2 | 1 | 2 | 1.3853 | 1.9 | < 2 | 5
| Mean |  | 1.5333 | 1.55 |  |  |  |  | 1.1441 | 2 | 1.1667 | 1.8333 | 1.4013 | 2.1333 |  | 4.5
| Median |  | 1.54 | 1.55 |  |  |  |  | 1.1481 | 2 | 1 | 2 | 1.3959 | 2.15 |  | 4.5
| Std Dev. |  | 0.02582 | 0.054772 |  |  |  |  | 0.027419 | 0 | 0.40825 | 0.40825 | 0.019752 | 0.2582 |  | 0.70711
| Rel.Std.Dev. |  | 1.684% | 3.534% |  |  |  |  | 2.397% | 0% | 34.99% | 22.27% | 1.41% | 12.1% |  | 15.71%
| PDM3 |  | 8.965% | 10.15% |  |  |  |  | -18.7% | 42.13% | -17.09% | 30.28% | -0.4167% | 51.6% |  | 219.8%
| Z-Score (Absolute) |  | 0.27 | 0.26 | 2.77 | 2.77 | 2.77 | 2.77 | 0.61 | NA | NA | NA | 0.39 | 0.26 | 0.74 | NA
Results from laboratories 9, 10, 11 and 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 353c (Indicative Value 0.044 ppm) | [15 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.09 | 1.10 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [10 empty] | 3
3
| 1 | 1 | 0.02 | 0.035022 | 0.06 | [11 empty] | 1
| 1 | 2 | 0.02 | 0.041904 | 0.06 | 0.05 | [10 empty] | 3
| 1 | 3 | 0.02 | 0.035222 | 0.06 | 0.05 | [10 empty] | 16
| 1 | 4 | 0.02 | 0.040452 | 0.07 | [11 empty] | 0.044311
| 1 | 5 | 0.02 | 0.039376 | 0.06 | [11 empty] | 9
| 1 | 6 | 0.02 | 0.051481 | 0.07 | 0.06
| Mean |  | 0.02 | 0.040576 | 0.063333 | 0.053333
| Median |  | 0.02 | 0.039914 | 0.06 | 0.05
| Std Dev. |  | 0 | 0.0060277 | 0.005164 | 0.0057735
| Rel.Std.Dev. |  | 0% | 14.86% | 8.154% | 10.83%
| PDM3 |  | -54.86% | -8.428% | 42.93% | 20.36%
| Z-Score (Absolute) |  | 1.6 | 0.38 | 0.97 | 0.38
| Analytical results for Mg in OREAS 353c (Certified Value 0.152 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.15 | 0.154 | 0.15 | 0.15 | 0.15 | 0.14 | 0.15332 | 0.15 | 0.16 | 0.16 | 0.15285 | 0.14 | 0.1605 | 0.16 | 1
| 1 | 2 | 0.16 | 0.147 | 0.15 | 0.14 | 0.15 | 0.14 | 0.15277 | 0.16 | 0.16 | 0.15 | 0.15346 | 0.14 | 0.1619 | 0.16 | #N/A
| 1 | 3 | 0.17 | 0.146 | 0.15 | 0.15 | 0.15 | 0.14 | 0.15306 | 0.16 | 0.16 | 0.15 | 0.15415 | 0.15 | 0.1585 | 0.16 | 16
| 1 | 4 | 0.16 | 0.151 | 0.15 | 0.14 | 0.14 | 0.14 | 0.155 | 0.15 | 0.16 | 0.15 | 0.1539 | 0.14 | 0.1601 | 0.16 | 0.15221
| 1 | 5 | 0.17 | 0.152 | 0.15 | 0.14 | 0.14 | 0.14 | 0.15383 | 0.15 | 0.16 | 0.15 | 0.1553 | 0.14 | 0.1638 | 0.16 | 29
| 1 | 6 | 0.15 | 0.152 | 0.15 | 0.15 | 0.15 | 0.14 | 0.15528 | 0.16 | 0.16 | 0.15 | 0.15471 | 0.15 | 0.1613 | 0.16
| Mean |  | 0.16 | 0.15033 | 0.15 | 0.145 | 0.14667 | 0.14 | 0.15388 | 0.155 | 0.16 | 0.15167 | 0.15406 | 0.14333 | 0.16102 | 0.16
| Median |  | 0.16 | 0.1515 | 0.15 | 0.145 | 0.15 | 0.14 | 0.15358 | 0.155 | 0.16 | 0.15 | 0.15403 | 0.14 | 0.1609 | 0.16
| Std Dev. |  | 0.0089443 | 0.0031411 | 0 | 0.0054772 | 0.005164 | 0 | 0.0010429 | 0.0054772 | 0 | 0.0040825 | 0.00087152 | 0.005164 | 0.0017916 | 0
| Rel.Std.Dev. |  | 5.59% | 2.089% | 0% | 3.777% | 3.521% | 0% | 0.6778% | 3.534% | 0% | 2.692% | 0.5657% | 3.603% | 1.113% | 0%
| PDM3 |  | 5.117% | -1.234% | -1.453% | -4.738% | -3.643% | -8.023% | 1.095% | 1.832% | 5.117% | -0.3578% | 1.217% | -5.833% | 5.785% | 5.117%
| Z-Score (Absolute) |  | 0.73 | 0.25 | 0.28 | 0.79 | 0.62 | 1.29 | 0.11 | 0.23 | 0.73 | 0.11 | 0.13 | 0.95 | 0.83 | 0.73
| Analytical results for Mn in OREAS 353c (Certified Value 0.136 wt.%) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.13863 | 0.128 | 0.137 | 0.135 | 0.131 | 0.136 | 0.12417 | 0.1309 | 0.1417 | 0.1409 | 0.13124 | 0.116 | 0.1501 | 0.1409 | 1
| 1 | 2 | 0.13754 | 0.124 | 0.142 | 0.133 | 0.1315 | 0.136 | 0.12408 | 0.1383 | 0.1419 | 0.1396 | 0.13195 | 0.117 | 0.1507 | 0.1421 | #N/A
| 1 | 3 | 0.13811 | 0.125 | 0.139 | 0.136 | 0.1325 | 0.1375 | 0.12399 | 0.1358 | 0.1412 | 0.1369 | 0.13185 | 0.122 | 0.1478 | 0.1409 | 16
| 1 | 4 | 0.13918 | 0.126 | 0.1405 | 0.1335 | 0.127 | 0.136 | 0.12239 | 0.1352 | 0.1427 | 0.1386 | 0.13245 | 0.113 | 0.15 | 0.141 | 0.136
| 1 | 5 | 0.13553 | 0.126 | 0.1375 | 0.1325 | 0.1305 | 0.1345 | 0.12221 | 0.1313 | 0.1415 | 0.1399 | 0.13339 | 0.117 | 0.1519 | 0.1404 | 30
| 1 | 6 | 0.14181 | 0.127 | 0.1415 | 0.1375 | 0.1325 | 0.134 | 0.12612 | 0.1359 | 0.142 | 0.1388 | 0.13386 | 0.125 | 0.1497 | 0.1409
| Mean |  | 0.13847 | 0.126 | 0.13958 | 0.13458 | 0.13083 | 0.13567 | 0.12383 | 0.13457 | 0.14183 | 0.13912 | 0.13246 | 0.11833 | 0.15003 | 0.14103
| Median |  | 0.13837 | 0.126 | 0.13975 | 0.13425 | 0.13125 | 0.136 | 0.12404 | 0.1355 | 0.1418 | 0.1392 | 0.1322 | 0.117 | 0.15005 | 0.1409
| Std Dev. |  | 0.0020639 | 0.0014142 | 0.0020837 | 0.0019343 | 0.0020412 | 0.0012517 | 0.0014245 | 0.0028898 | 0.00051251 | 0.0013644 | 0.00099685 | 0.0043665 | 0.0013441 | 0.00056451
| Rel.Std.Dev. |  | 1.491% | 1.122% | 1.493% | 1.437% | 1.56% | 0.9226% | 1.15% | 2.147% | 0.3613% | 0.9808% | 0.7526% | 3.69% | 0.8959% | 0.4003%
| PDM3 |  | 1.814% | -7.353% | 2.635% | -1.042% | -3.799% | -0.2451% | -8.951% | -1.054% | 4.289% | 2.292% | -2.605% | -12.99% | 10.32% | 3.701%
| Z-Score (Absolute) |  | 0.52 | 1.41 | 0.69 | 0.08 | 0.66 | 0.08 | 1.74 | 0.09 | 1.03 | 0.62 | 0.41 | 2.59 | 2.3 | 0.91
| Analytical results for Mo in OREAS 353c (Certified Value 9.95 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
2
| 1 | 1 | 7.34 | 9.7 | 11 | 11 | 11 | 10 | 8.92 | 7.27 | 11 | 10 | 9.6662 | 7.51 | 12 | 8 | 1
| 1 | 2 | 7.25 | 9.9 | 11 | 10 | 11 | 11 | 9.07 | 8.12 | 11 | 11 | 9.9788 | 7.82 | 12.2 | 9 | #N/A
| 1 | 3 | 6.63 | 10.1 | 11 | 11 | 11 | 11 | 8.66 | 7.74 | 11 | 10 | 9.9673 | 7.95 | 11.9 | 11 | 16
| 1 | 4 | 6.97 | 9.9 | 11 | 10 | 10 | 11 | 8.94 | 8.09 | 10 | 11 | 9.6244 | 7.68 | 12 | 8 | 9.9511
| 1 | 5 | 7.34 | 10 | 11 | 10 | 11 | 10 | 9.02 | 7.81 | 10 | 10 | 9.8971 | 7.98 | 12.2 | 10 | 31
| 1 | 6 | 6.93 | 9.8 | 11 | 11 | 10 | 10 | 9.03 | 7.51 | 11 | 10 | 9.8443 | 8.29 | 12.1 | 10
| Mean |  | 7.0767 | 9.9 | 11 | 10.5 | 10.6667 | 10.5 | 8.94 | 7.7567 | 10.6667 | 10.3333 | 9.8297 | 7.8717 | 12.0667 | 9.3333
| Median |  | 7.11 | 9.9 | 11 | 10.5 | 11 | 10.5 | 8.98 | 7.775 | 11 | 10 | 9.8707 | 7.885 | 12.05 | 9.5
| Std Dev. |  | 0.28324 | 0.14142 | 0 | 0.54772 | 0.5164 | 0.54772 | 0.14846 | 0.32983 | 0.5164 | 0.5164 | 0.15152 | 0.26947 | 0.12111 | 1.2111
| Rel.Std.Dev. |  | 4.002% | 1.428% | 0% | 5.216% | 4.841% | 5.216% | 1.661% | 4.252% | 4.841% | 4.997% | 1.541% | 3.423% | 1.004% | 12.98%
| PDM3 |  | -28.89% | -0.5138% | 10.54% | 5.516% | 7.191% | 5.516% | -10.16% | -22.05% | 7.191% | 3.841% | -1.22% | -20.9% | 21.26% | -6.208%
| Z-Score (Absolute) |  | 3.07 | 0.22 | 0.89 | 0.39 | 0.56 | 0.39 | 1.19 | 2.39 | 0.56 | 0.22 | 0.29 | 2.27 | 1.97 | 0.79
| Analytical results for Na in OREAS 353c (Indicative Value 0.014 wt.%) | [15 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.02 | 0.014 | 0.01 | 0.01 | 0.02 | 0.01 | 0.02 | 0.02 | 0.04 | 0.01319 | 0.02 | 0.0102 |  |  | 1
| 1 | 2 | 0.02 | 0.013 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.04 | 0.014147 | 0.01 | 0.0112 |  |  | 4
| 1 | 3 | 0.02 | 0.013 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.04 | 0.014541 | 0.02 | 0.0102 |  |  | 16
| 1 | 4 | 0.02 | 0.013 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 |  | 0.04 | 0.013422 | 0.01 | 0.0103 |  |  | 0.014014
| 1 | 5 | 0.02 | 0.013 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 |  | 0.04 | 0.01413 | 0.02 | 0.0107 |  |  | 10
| 1 | 6 | 0.02 | 0.013 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 |  | 0.04 | 0.013478 | 0.01 | 0.0104
| Mean |  | 0.02 | 0.013167 | 0.01 | 0.01 | 0.011667 | 0.01 | 0.02 | 0.02 | 0.04 | 0.013818 | 0.015 | 0.0105
| Median |  | 0.02 | 0.013 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.04 | 0.013804 | 0.015 | 0.01035
| Std Dev. |  | 0 | 0.00040825 | 0 | 0 | 0.0040825 | 0 | 0 | 0 | 0 | 0.00052833 | 0.0054772 | 0.00038987
| Rel.Std.Dev. |  | 0% | 3.101% | 0% | 0% | 34.99% | 0% | 0% | 0% | 0% | 3.823% | 36.51% | 3.713%
| PDM3 |  | 42.72% | -6.045% | -28.64% | -28.64% | -16.75% | -28.64% | 42.72% | 42.72% | 185.4% | -1.396% | 7.038% | -25.07%
| Z-Score (Absolute) |  | 1.26 | 0.06 | 0.67 | 0.67 | 0.35 | 0.67 | 1.26 | 1.26 | 5.12 | 0.06 | 0.29 | 0.58
| Analytical results for Nb in OREAS 353c (Indicative Value 0.26 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.09 | 1.12 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.32 | 0.2 | 0.27867 | 0.1 | 0.42 | [9 empty] | 1
| 1 | 2 | 0.32 | 0.2 | 0.28028 | 0.1 | 0.36 | [9 empty] | 5
| 1 | 3 | 0.34 | 0.2 | 0.27615 | 0.1 | 0.39 | [9 empty] | 16
| 1 | 4 | 0.34 | 0.2 | 0.27412 | < 0.1 | 0.4 | [9 empty] | 0.25738
| 1 | 5 | 0.33 | 0.2 | 0.28477 | 0.1 | 0.38 | [9 empty] | 11
| 1 | 6 | 0.31 | 0.2 | 0.28755 | 0.1 | 0.38
| Mean |  | 0.32667 | 0.2 | 0.28026 | 0.1 | 0.38833
| Median |  | 0.325 | 0.2 | 0.27947 | 0.1 | 0.385
| Std Dev. |  | 0.012111 | 3.0405e-17 | 0.0051107 | 0 | 0.020412
| Rel.Std.Dev. |  | 3.707% | 1.52e-14% | 1.824% | 0% | 5.256%
| PDM3 |  | 26.92% | -22.3% | 8.886% | -61.15% | 50.88%
| Z-Score (Absolute) |  | 0.39 | 0.67 | 0 | 1.58 | 0.91
| Analytical results for Nd in OREAS 353c (Indicative Value 6.44 ppm) | [15 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 5.57 | 5.1015 | 7.6 | 8.3 | [10 empty] | 1
| 1 | 2 | 5.37 | 5.21 | 6.5 | 8 | [10 empty] | 6
| 1 | 3 | 5.55 | 5.2476 | 7.1 | 8 | [10 empty] | 16
| 1 | 4 | 5.37 | 5.3135 | 7.1 | 7.8 | [10 empty] | 6.4434
| 1 | 5 | 5.35 | 5.1889 | 7.4 | 8.3 | [10 empty] | 12
| 1 | 6 | 5.43 | 5.0392 | 7 | 7.8
| Mean |  | 5.44 | 5.1834 | 7.1167 | 8.0333
| Median |  | 5.4 | 5.1995 | 7.1 | 8
| Std Dev. |  | 0.096954 | 0.099282 | 0.37639 | 0.22509
| Rel.Std.Dev. |  | 1.782% | 1.915% | 5.289% | 2.802%
| PDM3 |  | -15.57% | -19.55% | 10.45% | 24.68%
| Z-Score (Absolute) |  | 0.58 | 0.76 | 0.58 | 1.22
| Analytical results for Ni in OREAS 353c (Certified Value 15.8 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
1
| 1 | 1 | 19.33 | 15 | 17 | 16 | 16 | 14 | 15.4405 | 14.5 | 19 | 19 | 14.8704 | 13.9 | 18.1 | 18 | 1
| 1 | 2 | 19.22 | 15 | 16 | 16 | 16 | 14 | 15.5035 | 15.4 | 16 | 15 | 13.6974 | 13.9 | 18.4 | 16 | 18
| 1 | 3 | 18.61 | 15 | 14 | 16 | 16 | 15 | 14.443 | 15.3 | 14 | 19 | 14.7383 | 14.6 | 18.1 | 16 | 16
| 1 | 4 | 19.03 | 15 | 16 | 14 | 12 | 16 | 14.3695 | 14.9 | 16 | 25 | 14.2568 | 13.6 | 18.2 | 21 | 15.8123
| 1 | 5 | 18.86 | 15 | 17 | 16 | 17 | 15 | 14.5585 | 14.2 | 14 | 25 | 13.6582 | 14.1 | 18.4 | 16 | 32
| 1 | 6 | 19.54 | 15 | 17 | 17 | 15 | 15 | 15.2515 | 15 | 17 | 23 | 14.7783 | 14.3 | 18.1 | 15
| Mean |  | 19.0983 | 15 | 16.1667 | 15.8333 | 15.3333 | 14.8333 | 14.9277 | 14.8833 | 16 | 21 | 14.3332 | 14.0667 | 18.2167 | 17
| Median |  | 19.125 | 15 | 16.5 | 16 | 16 | 15 | 14.905 | 14.95 | 16 | 21 | 14.4976 | 14 | 18.15 | 16
| Std Dev. |  | 0.3357 | 0 | 1.169 | 0.98319 | 1.7512 | 0.75277 | 0.52577 | 0.46224 | 1.8974 | 4 | 0.55074 | 0.35024 | 0.1472 | 2.1909
| Rel.Std.Dev. |  | 1.758% | 0% | 7.231% | 6.21% | 11.42% | 5.075% | 3.522% | 3.106% | 11.86% | 19.05% | 3.842% | 2.49% | 0.808% | 12.89%
| PDM3 |  | 20.78% | -5.137% | 2.241% | 0.1333% | -3.029% | -6.191% | -5.594% | -5.875% | 1.187% | 32.81% | -9.354% | -11.04% | 15.21% | 7.512%
| Z-Score (Absolute) |  | 3.27 | 0.54 | 0.54 | 0.23 | 0.23 | 0.7 | 0.61 | 0.65 | 0.39 | 5.04 | 1.16 | 1.41 | 2.45 | 1.32
| Analytical results for P in OREAS 353c (Certified Value 0.013 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0158 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.044842 | 0.02 | 0.02 | 0.01 | 0.012893 | 0.0133 | 0.0138 | 0.01 | 1
| 1 | 2 | 0.0156 | 0.014 | 0.015 | 0.014 | 0.014 | 0.014 | 0.04434 | 0.02 | 0.02 | 0.02 | 0.013375 | 0.0122 | 0.014 | 0.01 | 19
| 1 | 3 | 0.0156 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.044203 | 0.02 | 0.02 | 0.01 | 0.013212 | 0.0136 | 0.0131 | 0.01 | 16
| 1 | 4 | 0.0156 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.04474 | 0.02 | 0.02 | 0.01 | 0.01314 | 0.0129 | 0.0144 | 0.01 | 0.013295
| 1 | 5 | 0.0155 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.045128 | 0.02 | 0.02 | 0.02 | 0.013694 | 0.0127 | 0.014 | 0.02 | 33
| 1 | 6 | 0.0155 | 0.014 | 0.015 | 0.014 | 0.014 | 0.014 | 0.044719 | 0.02 | 0.02 | 0.01 | 0.013641 | 0.0138 | 0.0141 | 0.01
| Mean |  | 0.0156 | 0.014 | 0.014333 | 0.014 | 0.014 | 0.014 | 0.044662 | 0.02 | 0.02 | 0.013333 | 0.013326 | 0.013083 | 0.0139 | 0.011667
| Median |  | 0.0156 | 0.014 | 0.014 | 0.014 | 0.014 | 0.014 | 0.044729 | 0.02 | 0.02 | 0.01 | 0.013294 | 0.0131 | 0.014 | 0.01
| Std Dev. |  | 0.00010954 | 0 | 0.0005164 | 0 | 0 | 0 | 0.0003386 | 0 | 0 | 0.005164 | 0.00030739 | 0.00059805 | 0.00043818 | 0.0040825
| Rel.Std.Dev. |  | 0.7022% | 0% | 3.603% | 0% | 0% | 0% | 0.7581% | 0% | 0% | 38.73% | 2.307% | 4.571% | 3.152% | 34.99%
| PDM3 |  | 17.34% | 5.305% | 7.812% | 5.305% | 5.305% | 5.305% | 235.9% | 50.44% | 50.44% | 0.2901% | 0.2329% | -1.59% | 4.552% | -12.25%
| Z-Score (Absolute) |  | 1.61 | 0 | 0.34 | 0 | 0 | 0 | 30.84 | 6.03 | 6.03 | 0.67 | 0.68 | 0.92 | 0.1 | 2.35
| Analytical results for Pb in OREAS 353c (Certified Value 70.09 wt.%) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.15 | [6 empty] | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*OES | EDTA*TITR | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | [6 empty] | 2
3
| 1 | 1 | 60.34 | 50.6 | 72.7 | 72.72 | > 70 | 66.3556 | 77.1258 | 72.2184 | [6 empty] | 1
| 1 | 2 | 60.71 | 51.1 | 71.4 | 72.63 | > 70 | 66.8098 | 78.5714 | 72.3804 | [6 empty] | 20
| 1 | 3 | 60.91 | 50.8 | 72.3 | 72.45 | > 70 | 68.2721 | 72.4855 | 72.1077 | [6 empty] | 16
| 1 | 4 | 60.8 | 51.3 | 71.1 | 72.49 | > 70 | 68.7752 | 71.2482 | 72.4208 | [6 empty] | 70.0908
| 1 | 5 | 61 | 50.7 | 71.4 | 72.71 | > 70 | 67.6953 | 80.8509 | 71.2881 | [6 empty] | 34
| 1 | 6 | 60.74 | 50.7 | 71.6 | 72.72 | > 70 | 67.8634 | 73.8029 | 72.2771
| Mean |  | 60.75 | 50.8667 | 71.75 | 72.62 |  | 67.6285 | 75.6808 | 72.1154
| Median |  | 60.77 | 50.75 | 71.5 | 72.67 |  | 67.7793 | 75.4643 | 72.2477
| Std Dev. |  | 0.22821 | 0.27325 | 0.61563 | 0.12166 |  | 0.90323 | 3.7565 | 0.42065
| Rel.Std.Dev. |  | 0.3757% | 0.5372% | 0.858% | 0.1675% |  | 1.336% | 4.964% | 0.5833%
| PDM3 |  | -13.33% | -27.43% | 2.367% | 3.608% |  | -3.513% | 7.975% | 2.889%
| Z-Score (Absolute) |  | 1.89 | 3.58 | 0 | 0.15 | NA | 0.71 | 0.67 | 0.06
| Analytical results for Pr in OREAS 353c (Indicative Value 1.76 ppm) | [15 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 1.47 | 1.3976 | 1.95 | 2.22 | [10 empty] | 1
| 1 | 2 | 1.4 | 1.5343 | 1.95 | 2.1 | [10 empty] | 7
| 1 | 3 | 1.43 | 1.6039 | 1.93 | 2.2 | [10 empty] | 16
| 1 | 4 | 1.49 | 1.5077 | 1.85 | 2.13 | [10 empty] | 1.7589
| 1 | 5 | 1.54 | 1.4705 | 1.95 | 2.16 | [10 empty] | 13
| 1 | 6 | 1.53 | 1.3734 | 1.8 | 2.1
| Mean |  | 1.4767 | 1.4812 | 1.905 | 2.1517
| Median |  | 1.48 | 1.4891 | 1.94 | 2.145
| Std Dev. |  | 0.055015 | 0.086377 | 0.06442 | 0.050761
| Rel.Std.Dev. |  | 3.726% | 5.831% | 3.382% | 2.359%
| PDM3 |  | -16.05% | -15.79% | 8.307% | 22.33%
| Z-Score (Absolute) |  | 0.68 | 0.67 | 0.67 | 1.44
| Analytical results for Rb in OREAS 353c (Indicative Value 1.27 ppm) | [15 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.75 | 1.1 | 0.81342 | 1.5 | 1 | 1.5 | 1.1385 | 1.4 | 1.3 | [5 empty] | 1
| 1 | 2 | 1.88 | 1.1 | 0.80135 | 2.7 | 1.3 | 1.5 | 1.2073 | 1.4 | 1.1 | [5 empty] | 8
| 1 | 3 | 1.71 | 1.1 | 0.70509 | 1.7 | 1 | 1.8 | 1.1871 | 1.4 | 1.2 | [5 empty] | 16
| 1 | 4 | 1.83 | 1.1 | 0.78146 | 1.8 | 1.5 | 1.6 | 1.0582 | 1.2 | 1.3 | [5 empty] | 1.2705
| 1 | 5 | 1.84 | 1.1 | 0.79107 | 1.7 | 1.2 | 1.7 | 1.211 | 1.5 | 1.2 | [5 empty] | 14
| 1 | 6 | 1.8 | 1.1 | 0.72895 | 1.8 | 1.4 | 1.6 | 1.0615 | 1.4 | 1.2
| Mean |  | 1.8017 | 1.1 | 0.77023 | 1.8667 | 1.2333 | 1.6167 | 1.1439 | 1.3833 | 1.2167
| Median |  | 1.815 | 1.1 | 0.78627 | 1.75 | 1.25 | 1.6 | 1.1628 | 1.4 | 1.2
| Std Dev. |  | 0.062423 | 2.4324e-16 | 0.043226 | 0.42269 | 0.20656 | 0.1169 | 0.070101 | 0.098319 | 0.075277
| Rel.Std.Dev. |  | 3.465% | 2.211e-14% | 5.612% | 22.64% | 16.75% | 7.231% | 6.128% | 7.107% | 6.187%
| PDM3 |  | 41.81% | -13.42% | -39.38% | 46.92% | -2.927% | 27.24% | -9.963% | 8.879% | -4.239%
| Z-Score (Absolute) |  | 2.55 | 0.6 | 2.08 | 2.85 | 0 | 1.72 | 0.4 | 0.67 | 0.07
| Analytical results for Re in OREAS 353c (Indicative Value < 0.05 ppm) | [15 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.12 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [9 empty] | 3
3
| 1 | 1 | 0.05 | < 0.05 | < 0.05 | < 0.002 | < 0.003 | [9 empty] | 1
| 1 | 2 | 0.05 | < 0.05 | < 0.05 | < 0.002 | < 0.003 | [9 empty] | 9
| 1 | 3 | 0.05 | < 0.05 | < 0.05 | < 0.002 | < 0.003 | [9 empty] | 16
| 1 | 4 | 0.05 | < 0.05 | < 0.05 | < 0.002 | < 0.003 | [9 empty] | < 0.05
| 1 | 5 | 0.05 | < 0.05 | < 0.05 | < 0.002 | 0.003 | [9 empty] | 15
| 1 | 6 | 0.05 | < 0.05 | < 0.05 | < 0.002 | < 0.003
| Mean |  | 0.05 |  |  |  | 0.003
| Median |  | 0.05 |  |  |  | 0.003
| Std Dev. |  | 7.6012e-18
| Rel.Std.Dev. |  | 1.52e-14%
PDM3
| Z-Score (Absolute) |  | 0.73 | 0 | 0 | 0.7 | 0.67
| Analytical results for S in OREAS 353c (Certified Value 14.23 wt.%) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 14.4 | 13.1 | 13.75 | 14.65 | 14.4 | 14.35 | 14.2274 | > 5 | 13.95 | 13.68 | 14.4817 | 12.9 | 13.3211 | 14.46 | 1
| 1 | 2 | 14.55 | 13.3 | 13.85 | 14.75 | 14.4 | 14.45 | 14.7191 | > 5 | 13.86 | 13.5 | 14.5138 | 12.8 | 14.4541 | 14.64 | #N/A
| 1 | 3 | 14.56 | 13.2 | 13.9 | 14.6 | 14.45 | 14.6 | 14.9597 | > 5 | 13.82 | 13.23 | 14.2101 | 12.9 | 13.8134 | 14.47 | 16
| 1 | 4 | 14.34 | 13.2 | 13.7 | 14.75 | 14.5 | 14.4 | 14.3443 | > 5 | 14.05 | 13.43 | 14.4311 | 12.5 | 14.2347 | 14.32 | 14.2259
| 1 | 5 | 14.4 | 13.1 | 13.7 | 14.7 | 14.4 | 14.35 | 14.5034 | > 5 | 13.71 | 13.57 | 14.0866 | 12.8 | 13.9903 | 14.46 | 35
| 1 | 6 | 14.7 | 12.9 | 13.7 | 14.2 | 14.35 | 14.55 | 14.6992 | > 5 | 14.13 | 13.42 | 14.2715 | 12.9 | 14.2285 | 14.32
| Mean |  | 14.4917 | 13.1333 | 13.7667 | 14.6083 | 14.4167 | 14.45 | 14.5755 |  | 13.92 | 13.4717 | 14.3325 | 12.8 | 14.007 | 14.445
| Median |  | 14.475 | 13.15 | 13.725 | 14.675 | 14.4 | 14.425 | 14.6013 |  | 13.905 | 13.465 | 14.3513 | 12.85 | 14.1094 | 14.46
| Std Dev. |  | 0.13512 | 0.13663 | 0.08756 | 0.20837 | 0.05164 | 0.10488 | 0.26958 |  | 0.15466 | 0.15276 | 0.16971 | 0.15492 | 0.40225 | 0.11862
| Rel.Std.Dev. |  | 0.9324% | 1.04% | 0.636% | 1.426% | 0.3582% | 0.7258% | 1.85% |  | 1.111% | 1.134% | 1.184% | 1.21% | 2.872% | 0.8212%
| PDM3 |  | 1.868% | -7.68% | -3.228% | 2.688% | 1.341% | 1.575% | 2.458% |  | -2.15% | -5.302% | 0.749% | -10.02% | -1.539% | 1.54%
| Z-Score (Absolute) |  | 0.39 | 2.93 | 1.38 | 0.67 | 0.21 | 0.29 | 0.59 | NA | 1.01 | 2.1 | 0 | 3.75 | 0.8 | 0.28
| Analytical results for Sb in OREAS 353c (Certified Value 0.266 wt.%) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.23066 | 0.258 | 0.275 | 0.289 | 0.284 | 0.28 | 0.24786 | 0.1482 | 0.2784 | 0.25622 | 0.25239 | 0.171 | 0.28333 | 0.28048 | 1
| 1 | 2 | 0.23438 | 0.253 | 0.284 | 0.283 | 0.279 | 0.284 | 0.24841 | 0.1742 | 0.28013 | 0.27423 | 0.25346 | 0.18 | 0.22445 | 0.28227 | #N/A
| 1 | 3 | 0.23041 | 0.254 | 0.276 | 0.286 | 0.282 | 0.283 | 0.24895 | 0.1655 | 0.27478 | 0.25283 | 0.25461 | 0.182 | 0.257 | 0.2783 | 16
| 1 | 4 | 0.23182 | 0.245 | 0.281 | 0.28 | 0.274 | 0.286 | 0.24711 | 0.1588 | 0.28256 | 0.27698 | 0.25532 | 0.181 | 0.2599 | 0.28537 | 0.26592
| 1 | 5 | 0.23324 | 0.254 | 0.275 | 0.279 | 0.274 | 0.28 | 0.24719 | 0.1625 | 0.27586 | 0.26484 | 0.25543 | 0.186 | 0.2492 | 0.28655 | 36
| 1 | 6 | 0.23118 | 0.247 | 0.283 | 0.289 | 0.284 | 0.279 | 0.24803 | 0.1356 | 0.27872 | 0.2659 | 0.25705 | 0.193 | 0.24927 | 0.28146
| Mean |  | 0.23195 | 0.25183 | 0.279 | 0.28433 | 0.2795 | 0.282 | 0.24792 | 0.15747 | 0.27841 | 0.26517 | 0.25471 | 0.18217 | 0.25386 | 0.28241
| Median |  | 0.2315 | 0.2535 | 0.2785 | 0.2845 | 0.2805 | 0.2815 | 0.24795 | 0.16065 | 0.27856 | 0.26537 | 0.25496 | 0.1815 | 0.25313 | 0.28186
| Std Dev. |  | 0.0015637 | 0.0048751 | 0.0041473 | 0.0043665 | 0.0046368 | 0.0027568 | 0.00070789 | 0.01368 | 0.0028269 | 0.009535 | 0.0016322 | 0.0072503 | 0.019092 | 0.0030803
| Rel.Std.Dev. |  | 0.6742% | 1.936% | 1.486% | 1.536% | 1.659% | 0.9776% | 0.2855% | 8.687% | 1.015% | 3.596% | 0.6408% | 3.98% | 7.521% | 1.091%
| PDM3 |  | -12.78% | -5.299% | 4.917% | 6.923% | 5.105% | 6.045% | -6.768% | -40.79% | 4.695% | -0.2847% | -4.217% | -31.5% | -4.537% | 6.198%
| Z-Score (Absolute) |  | 0.98 | 0.28 | 0.67 | 0.85 | 0.68 | 0.77 | 0.42 | 3.58 | 0.64 | 0.18 | 0.18 | 2.72 | 0.21 | 0.78
| Analytical results for Sc in OREAS 353c (Certified Value < 1 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.12 | 1.13 | 1.14 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 0.25 | < 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.18 |  |  |  |  | 1
| 1 | 2 | 0.25 | < 1 | < 1 | < 1 | < 1 | < 1 | 0.4 | < 1 | 0.2 | 0.19 |  |  |  |  | 24
| 1 | 3 | 0.24 | < 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.18 |  |  |  |  | 16
| 1 | 4 | 0.27 | < 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.18 |  |  |  |  | < 1
| 1 | 5 | 0.27 | < 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.18 |  |  |  |  | 37
| 1 | 6 | 0.26 | < 1 | < 1 | < 1 | < 1 | < 1 | 0.3 | < 1 | 0.2 | 0.18
| Mean |  | 0.25667 | [5 empty] | 0.31667 |  | 0.2 | 0.18167
| Median |  | 0.255 | [5 empty] | 0.3 |  | 0.2 | 0.18
| Std Dev. |  | 0.012111 | [5 empty] | 0.040825 |  | 3.0405e-17 | 0.0040825
| Rel.Std.Dev. |  | 4.718% | [5 empty] | 12.89% |  | 1.52e-14% | 2.247%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Se in OREAS 353c (Indicative Value 5.16 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | [5 empty] | 2
2
| 1 | 1 | 4.73 | 6 | 6 | 6 | 6 | 5.0285 | 3.6 | 4.22 | [6 empty] | 1
| 1 | 2 | 4.37 | 6 | 8 | 5 |  | 4.8462 | 3.6 | 3.37 | [6 empty] | 10
| 1 | 3 | 4.21 | 6 | 7 |  |  | 4.3344 | 3.6 | 3.99 | 5 | [5 empty] | 16
| 1 | 4 | 4.46 | 6 | 8 | 6 | 6 | 4.5393 | 3.4 | 4.06 | [6 empty] | 5.1579
| 1 | 5 | 4.38 | 6 | 8 | 6 |  | 5.7042 | 3.6 | 3.71 | [6 empty] | 16
| 1 | 6 | 4.51 | 6 | 4 |  |  | 5.6118 | 3.9 | 3.59
| Mean |  | 4.4433 | 6 | 6.8333 | 5.75 | 6 | 5.0107 | 3.6167 | 3.8233 | 5
| Median |  | 4.42 | 6 | 7.5 | 6 | 6 | 4.9373 | 3.6 | 3.85 | 5
| Std Dev. |  | 0.17363 | 0 | 1.6021 | 0.5 | 0 | 0.55665 | 0.16021 | 0.3206
| Rel.Std.Dev. |  | 3.908% | 0% | 23.45% | 8.696% | 0% | 11.11% | 4.43% | 8.385%
| PDM3 |  | -13.85% | 16.33% | 32.48% | 11.48% | 16.33% | -2.853% | -29.88% | -25.87% | -3.061%
| Z-Score (Absolute) |  | 0.39 | 0.67 | 1.24 | 0.5 | 0.67 | 0 | 0.95 | 0.81 | 0.01
| Analytical results for Sm in OREAS 353c (Indicative Value 1.25 ppm) | [15 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.08 | 1.10 | 1.11 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
2
| 1 | 1 | 0.9 | 0.98697 | 1.4 | 2.1 | [10 empty] | 1
| 1 | 2 | 0.9 | 1.0001 | 1.3 | 1.7 | [10 empty] | 11
| 1 | 3 | 0.87 | 0.98405 | 1.7 | 1.7 | [10 empty] | 16
| 1 | 4 | 0.91 | 1.0444 | 1.3 | 1.7 | [10 empty] | 1.2514
| 1 | 5 | 0.89 | 0.96478 | 1.3 | 1.7 | [10 empty] | 17
| 1 | 6 | 0.89 | 0.99448 | 1.2 | 1.6
| Mean |  | 0.89333 | 0.9958 | 1.3667 | 1.75
| Median |  | 0.895 | 0.99073 | 1.3 | 1.7
| Std Dev. |  | 0.013663 | 0.026693 | 0.17512 | 0.17607
| Rel.Std.Dev. |  | 1.529% | 2.681% | 12.81% | 10.06%
| PDM3 |  | -28.62% | -20.43% | 9.207% | 39.84%
| Z-Score (Absolute) |  | 0.82 | 0.53 | 0.53 | 1.62
| Analytical results for Sn in OREAS 353c (Certified Value 12.5 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 1
1
| 1 | 1 | 7.39 | 13.2 | 10.8 | 13 | 14 | 12.1472 | 11.2 | 16.43 | [6 empty] | 1
| 1 | 2 | 7.82 | 13.4 | 11.7 | 15 | 15 | 12.7245 | 11.8 | 13.3 | [6 empty] | 14
| 1 | 3 | 7.75 | 13.2 | 11.8 | 15 | 14 | 12.6657 | 12 | 14.81 | [6 empty] | 16
| 1 | 4 | 7.47 | 13.3 | 12 | 15 | 14 | 12.4028 | 11.3 | 15.2 | [6 empty] | 12.5212
| 1 | 5 | 7.81 | 13.1 | 11.5 | 15 | 15 | 12.913 | 11.9 | 15.44 | [6 empty] | 38
| 1 | 6 | 7.61 | 13.2 | 10.6 | 14 | 14 | 12.3553 | 12 | 13.78
| Mean |  | 7.6417 | 13.2333 | 11.4 | 14.5 | 14.3333 | 12.5347 | 11.7 | 14.8267
| Median |  | 7.68 | 13.2 | 11.6 | 15 | 14 | 12.5343 | 11.85 | 15.005
| Std Dev. |  | 0.18203 | 0.10328 | 0.56921 | 0.83666 | 0.5164 | 0.28144 | 0.35777 | 1.1413
| Rel.Std.Dev. |  | 2.382% | 0.7805% | 4.993% | 5.77% | 3.603% | 2.245% | 3.058% | 7.698%
| PDM3 |  | -38.97% | 5.687% | -8.955% | 15.8% | 14.47% | 0.108% | -6.559% | 18.41%
| Z-Score (Absolute) |  | 2.41 | 0.16 | 0.68 | 0.74 | 0.67 | 0.16 | 0.54 | 0.89
| Analytical results for Sr in OREAS 353c (Indicative Value 2.15 ppm) | [15 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 2.15 | 1.9 | 1 | 2 | 3 | 3 | 2.9 | 1.6 | 2.3 | 2.5277 | 2 | < 5 |  |  | 1
| 1 | 2 | 2.07 | 1.8 | 2 | 3 | 3 | 5 | 6.1 | 1.8 | 2.3 | 2.5206 | 1.9 | < 5 |  |  | 3
| 1 | 3 | 2.1 | 1.9 | 1 | 2 | 3 | 3 | 3.9 | 2 | 2.3 | 2.2782 | 2.2 | < 5 |  |  | 16
| 1 | 4 | 2.13 | 1.8 | 1 | 2 | 3 | 3 | 3.5 | 1.8 | 2.4 | 2.1393 | 1.8 | < 5 |  |  | 2.1489
| 1 | 5 | 2.08 | 1.9 | 1 | 2 | 2 | 4 | 3.3 | 1.8 | 2.4 | 2.26 | 2.4 | < 5 |  |  | 9
| 1 | 6 | 2.02 | 1.8 | 1 | 2 | 2 | 2 | 3.4 | 2 | 2.5 | 2.1555 | 1.9 | < 5
| Mean |  | 2.0917 | 1.85 | 1.1667 | 2.1667 | 2.6667 | 3.3333 | 3.85 | 1.8333 | 2.3667 | 2.3136 | 2.0333
| Median |  | 2.09 | 1.85 | 1 | 2 | 3 | 3 | 3.45 | 1.8 | 2.35 | 2.2691 | 1.95
| Std Dev. |  | 0.046224 | 0.054772 | 0.40825 | 0.40825 | 0.5164 | 1.0328 | 1.1485 | 0.15055 | 0.08165 | 0.17215 | 0.22509
| Rel.Std.Dev. |  | 2.21% | 2.961% | 34.99% | 18.84% | 19.36% | 30.98% | 29.83% | 8.212% | 3.45% | 7.441% | 11.07%
| PDM3 |  | -2.661% | -13.91% | -45.71% | 0.8289% | 24.1% | 55.12% | 79.17% | -14.68% | 10.14% | 7.664% | -5.376%
| Z-Score (Absolute) |  | 0.31 | 0.81 | 2.23 | 0.15 | 0.89 | 2.27 | 3.34 | 0.84 | 0.26 | 0.15 | 0.43 | 0.54
| Analytical results for Ta in OREAS 353c (Certified Value < 0.05 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.05 | < 0.05 | 0.1 | < 0.05 | < 0.05 | 0.23 | [7 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.05 |  | < 0.05 | < 0.05 | < 0.03 | [7 empty] | 15
| 1 | 3 | < 0.05 | < 0.05 | < 0.05 |  | < 0.05 | < 0.05 | < 0.03 | [7 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.05 |  | < 0.05 | < 0.05 | 0.28 | [7 empty] | < 0.05
| 1 | 5 | < 0.05 | < 0.05 | < 0.05 |  | < 0.05 | < 0.05 | 0.07 | [7 empty] | 39
| 1 | 6 | < 0.05 | < 0.05 | < 0.05 |  | < 0.05 | < 0.05 | < 0.03
| Mean |  |  |  |  | 0.1 |  |  | 0.19333
| Median |  |  |  |  | 0.1 |  |  | 0.23
| Std Dev. | [7 empty] | 0.1097
| Rel.Std.Dev. | [7 empty] | 56.74%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 353c (Indicative Value 0.16 ppm) | [15 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.08 | 1.09 | 1.10 | 1.11 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 0.08 | 0.13288 | 0.23 | 0.14 | 0.19 | [9 empty] | 1
| 1 | 2 | 0.08 | 0.13026 | 0.23 | 0.14 | 0.2 | [9 empty] | 4
| 1 | 3 | 0.08 | 0.13577 | 0.23 | 0.19 | 0.21 | [9 empty] | 16
| 1 | 4 | 0.08 | 0.13527 | 0.24 | 0.12 | 0.17 | [9 empty] | 0.15753
| 1 | 5 | 0.09 | 0.1194 | 0.22 | 0.17 | 0.19 | [9 empty] | 10
| 1 | 6 | 0.08 | 0.14242 | 0.24 | 0.14 | 0.19
| Mean |  | 0.081667 | 0.13267 | 0.23167 | 0.15 | 0.19167
| Median |  | 0.08 | 0.13407 | 0.23 | 0.14 | 0.19
| Std Dev. |  | 0.0040825 | 0.0076609 | 0.0075277 | 0.025298 | 0.013292
| Rel.Std.Dev. |  | 4.999% | 5.775% | 3.249% | 16.87% | 6.935%
| PDM3 |  | -48.16% | -15.78% | 47.06% | -4.782% | 21.67%
| Z-Score (Absolute) |  | 1.11 | 0.28 | 1.32 | 0 | 0.67
| Analytical results for Te in OREAS 353c (Indicative Value 0.69 ppm) | [15 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | [5 empty] | 2
2
| 1 | 1 | 0.43 | 0.7 | 0.46 | 0.7 | 0.7 | 0.68459 | 0.52 | 0.89 | 0.8 | [5 empty] | 1
| 1 | 2 | 0.41 | 0.8 | 0.58 | 1.2 | 0.6 | 0.707 | 0.58 | 0.66 | 0.9 | [5 empty] | 5
| 1 | 3 | 0.43 | 0.8 | 0.54 | 1.2 | 0.8 | 0.70198 | 0.54 | 0.8 | 1 | [5 empty] | 16
| 1 | 4 | 0.41 | 0.7 | 0.49 | 1 | 1.1 | 0.63948 | 0.5 | 0.88 | 0.8 | [5 empty] | 0.69187
| 1 | 5 | 0.44 | 0.7 | 0.53 | 0.7 | 1 | 0.59691 | 0.45 | 0.89 | 0.7 | [5 empty] | 11
| 1 | 6 | 0.41 | 0.7 | 0.45 | 0.6 | 0.9 | 0.63099 | 0.54 | 0.57 | 0.9
| Mean |  | 0.42167 | 0.73333 | 0.50833 | 0.9 | 0.85 | 0.66016 | 0.52167 | 0.78167 | 0.85
| Median |  | 0.42 | 0.7 | 0.51 | 0.85 | 0.85 | 0.66203 | 0.53 | 0.84 | 0.85
| Std Dev. |  | 0.013292 | 0.05164 | 0.050365 | 0.26833 | 0.18708 | 0.044315 | 0.044008 | 0.13644 | 0.10488
| Rel.Std.Dev. |  | 3.152% | 7.042% | 9.908% | 29.81% | 22.01% | 6.713% | 8.436% | 17.46% | 12.34%
| PDM3 |  | -39.05% | 5.993% | -26.53% | 30.08% | 22.86% | -4.583% | -24.6% | 12.98% | 22.86%
| Z-Score (Absolute) |  | 1.8 | 0 | 1.3 | 0.96 | 0.67 | 0.42 | 1.22 | 0.28 | 0.67
| Analytical results for Th in OREAS 353c (Certified Value 0.41 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.5 | 0.45 | < 20 | < 20 | < 20 | < 20 | 0.4 | 0.4 | 0.4 | 0.36543 | 0.4 | 0.4 |  |  | 1
| 1 | 2 | 0.53 | 0.43 | < 20 | < 20 | < 20 | < 20 | 0.6 | 0.4 | 0.4 | 0.37155 | 0.4 | 0.4 |  |  | 18
| 1 | 3 | 0.5 | 0.44 | < 20 | < 20 | < 20 | < 20 | 0.5 | 0.4 | 0.4 | 0.38327 | 0.4 | 0.38 |  |  | 16
| 1 | 4 | 0.5 | 0.45 | < 20 | < 20 | < 20 | < 20 | 0.5 | 0.4 | 0.4 | 0.35784 | 0.4 | 0.41 |  |  | 0.41286
| 1 | 5 | 0.49 | 0.46 | < 20 | < 20 | < 20 | < 20 | 0.5 | 0.4 | 0.5 | 0.38195 | 0.4 | 0.37 |  |  | 40
| 1 | 6 | 0.5 | 0.45 | < 20 | < 20 | < 20 | < 20 | 0.4 | 0.4 | 0.4 | 0.35014 | 0.4 | 0.43
| Mean |  | 0.50333 | 0.44667 |  |  |  |  | 0.48333 | 0.4 | 0.41667 | 0.36836 | 0.4 | 0.39833
| Median |  | 0.5 | 0.45 |  |  |  |  | 0.5 | 0.4 | 0.4 | 0.36849 | 0.4 | 0.4
| Std Dev. |  | 0.013663 | 0.010328 |  |  |  |  | 0.075277 | 6.0809e-17 | 0.040825 | 0.013178 | 6.0809e-17 | 0.02137
| Rel.Std.Dev. |  | 2.714% | 2.312% |  |  |  |  | 15.57% | 1.52e-14% | 9.798% | 3.577% | 1.52e-14% | 5.365%
| PDM3 |  | 21.91% | 8.188% |  |  |  |  | 17.07% | -3.115% | 0.9217% | -10.78% | -3.115% | -3.519%
| Z-Score (Absolute) |  | 0.39 | 0.19 | 97.66 | 97.66 | 97.66 | 97.66 | 0.19 | 0.67 | 0.5 | 0.99 | 0.67 | 0.68
| Analytical results for Ti in OREAS 353c (Indicative Value 0.01 wt.%) | [15 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.01 | 0.0071 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.0056 | 0.009 | 0.0056624 | < 0.01 | 0.0062 |  |  | 1
| 1 | 2 | 0.01 | 0.0067 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.0057 | 0.0088 | 0.0055889 | < 0.01 | 0.0062 |  |  | 6
| 1 | 3 | 0.01 | 0.0067 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.0056 | 0.0094 | 0.0058606 | < 0.01 | 0.006 |  |  | 16
| 1 | 4 | 0.01 | 0.0067 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.0058 | 0.0081 | 0.0058945 | < 0.01 | 0.006 |  |  | 0.0098133
| 1 | 5 | 0.01 | 0.0066 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.0054 | 0.0089 | 0.0055806 | < 0.01 | 0.0061 |  |  | 12
| 1 | 6 | 0.01 | 0.0069 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.0056 | 0.0083 | 0.0054045 | < 0.01 | 0.006
| Mean |  | 0.01 | 0.0067833 | 0.01 | 0.01 | 0.01 | 0.01 |  | 0.0056167 | 0.00875 | 0.0056653 |  | 0.0060833
| Median |  | 0.01 | 0.0067 | 0.01 | 0.01 | 0.01 | 0.01 |  | 0.0056 | 0.00885 | 0.0056257 |  | 0.00605
| Std Dev. |  | 0 | 0.00018348 | 0 | 0 | 0 | 0 |  | 0.00013292 | 0.00047645 | 0.00018532 |  | 9.8319e-05
| Rel.Std.Dev. |  | 0% | 2.705% | 0% | 0% | 0% | 0% |  | 2.366% | 5.445% | 3.271% |  | 1.616%
| PDM3 |  | 1.902% | -30.88% | 1.902% | 1.902% | 1.902% | 1.902% |  | -42.76% | -10.84% | -42.27% |  | -38.01%
| Z-Score (Absolute) |  | 0.67 | 0.3 | 0.67 | 0.67 | 0.67 | 0.67 | 0.84 | 0.65 | 0.3 | 0.63 | 0.84 | 0.51
| Analytical results for Tl in OREAS 353c (Certified Value 1.92 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.81 | 2.19 | < 10 | < 10 | < 10 | < 10 | 1.95 | 2 | 3.3 | 1.909 | 1.68 | 1.766 | 2.1 |  | 1
| 1 | 2 | 1.83 | 2.14 | < 10 | < 10 | < 10 | < 10 | 2.05 | 2.1 | 3.1 | 2.0011 | 1.64 | 1.637 | 1.9 |  | #N/A
| 1 | 3 | 1.8 | 2.12 | < 10 | 10 | < 10 | < 10 | 2.01 | 2 | 3.3 | 1.9941 | 1.79 | 1.641 | 1.9 |  | 16
| 1 | 4 | 1.84 | 2.1 | < 10 | < 10 | < 10 | < 10 | 2.06 | 2 | 3.3 | 1.9528 | 1.67 | 1.772 | 1.9 |  | 1.9218
| 1 | 5 | 1.82 | 2.11 | < 10 | < 10 | < 10 | < 10 | 1.96 | 2.1 | 3.2 | 1.942 | 1.76 | 2.021 | 1.9 |  | 41
| 1 | 6 | 1.77 | 2.14 | < 10 | < 10 | < 10 | < 10 | 2.06 | 1.9 | 3.3 | 1.9368 | 1.77 | 1.905 | 1.9
| Mean |  | 1.8117 | 2.1333 |  | 10 |  |  | 2.015 | 2.0167 | 3.25 | 1.956 | 1.7183 | 1.7903 | 1.9333
| Median |  | 1.815 | 2.13 |  | 10 |  |  | 2.03 | 2 | 3.3 | 1.9474 | 1.72 | 1.769 | 1.9
| Std Dev. |  | 0.024833 | 0.032042 |  |  |  |  | 0.0501 | 0.075277 | 0.083666 | 0.035399 | 0.062423 | 0.1504 | 0.08165
| Rel.Std.Dev. |  | 1.371% | 1.502% |  |  |  |  | 2.486% | 3.733% | 2.574% | 1.81% | 3.633% | 8.4% | 4.223%
| PDM3 |  | -5.732% | 11.01% |  | 420.3% |  |  | 4.848% | 4.935% | 69.11% | 1.777% | -10.59% | -6.842% | 0.5985%
| Z-Score (Absolute) |  | 0.61 | 0.35 | 8.89 | 11.37 | 8.89 | 8.89 | 0 | 0 | 3.67 | 0.18 | 0.89 | 0.67 | 0.25
| Analytical results for Tm in OREAS 353c (Indicative Value 0.045 ppm) | [15 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.08 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | [12 empty] | 3
3
| 1 | 1 | 0.04 | 0.042173 | [12 empty] | 1
| 1 | 2 | 0.05 | 0.046311 | [12 empty] | 7
| 1 | 3 | 0.05 | 0.045234 | [12 empty] | 16
| 1 | 4 | 0.05 | 0.04779 | [12 empty] | 0.045447
| 1 | 5 | 0.04 | 0.042523 | [12 empty] | 13
| 1 | 6 | 0.05 | 0.041337
| Mean |  | 0.046667 | 0.044228
| Median |  | 0.05 | 0.043878
| Std Dev. |  | 0.005164 | 0.0025894
| Rel.Std.Dev. |  | 11.07% | 5.855%
| PDM3 |  | 2.683% | -2.683%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 353c (Certified Value 1.05 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.96 | 1.26 | < 10 | < 10 | < 10 | < 10 | 0.64134 | 0.96 | 1 | 1 | 1.058 | 0.961 | 0.91 |  | 1
| 1 | 2 | 1.02 | 1.2 | < 10 | < 10 | < 10 | < 10 | 0.63572 | 1.01 | 1.1 | 1.1 | 1.093 | 0.955 | 0.89 |  | #N/A
| 1 | 3 | 1.03 | 1.22 | < 10 | < 10 | < 10 | < 10 | 0.6429 | 1 | 1.2 | 1.1 | 1.078 | 1.02 | 0.86 |  | 16
| 1 | 4 | 0.99 | 1.26 | < 10 | < 10 | < 10 | < 10 | 0.67715 | 1.02 | 1 | 1.1 | 1.0575 | 0.94 | 0.92 |  | 1.0488
| 1 | 5 | 1.02 | 1.23 | < 10 | < 10 | < 10 | < 10 | 0.63736 | 0.99 | 1.2 | 1.1 | 1.0796 | 1.01 | 0.87 |  | 42
| 1 | 6 | 1.04 | 1.26 | < 10 | < 10 | < 10 | < 10 | 0.64395 | 1.03 | 1.1 | 1 | 1.0794 | 1.05 | 1.01
| Mean |  | 1.01 | 1.2383 |  |  |  |  | 0.6464 | 1.0017 | 1.1 | 1.0667 | 1.0743 | 0.98933 | 0.91
| Median |  | 1.02 | 1.245 |  |  |  |  | 0.64212 | 1.005 | 1.1 | 1.1 | 1.0787 | 0.9855 | 0.9
| Std Dev. |  | 0.029665 | 0.025626 |  |  |  |  | 0.015397 | 0.024833 | 0.089443 | 0.05164 | 0.013897 | 0.043505 | 0.054037
| Rel.Std.Dev. |  | 2.937% | 2.069% |  |  |  |  | 2.382% | 2.479% | 8.131% | 4.841% | 1.294% | 4.397% | 5.938%
| PDM3 |  | -3.698% | 18.07% |  |  |  |  | -38.37% | -4.492% | 4.884% | 1.705% | 2.428% | -5.668% | -13.23%
| Z-Score (Absolute) |  | 0.26 | 0.67 | 16.13 | 16.13 | 16.13 | 16.13 | 1.76 | 0.3 | 0.11 | 0.03 | 0 | 0.35 | 0.67
| Analytical results for V in OREAS 353c (Certified Value 5.02 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.12 | 1.13 | 1.14 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 4.49 | 6 | 5 | 5 | 5 | 5 | 5.2128 | 4 | 4.3995 | 3.5 | 6 |  |  |  | 1
| 1 | 2 | 4.75 | 6 | 5 | 5 | 5 | 5 | 5.105 | 4 | 5.1083 | 3.6 | 6 |  |  |  | 21
| 1 | 3 | 4.65 | 6 | 5 | 5 | 5 | 5 | 5.2422 | 4 | 4.1917 | 3.7 | 6 |  |  |  | 16
| 1 | 4 | 4.68 | 6 | 5 | 5 | 5 | 5 | 5.0364 | 4 | 4.9045 | 3.6 | 6 |  |  |  | 5.025
| 1 | 5 | 4.63 | 6 | 5 | 5 | 5 | 5 | 4.8992 | 4 | 3.9435 | 3.7 | 6 |  |  |  | 43
| 1 | 6 | 4.71 | 6 | 5 | 5 | 5 | 5 | 5.2772 | 4 | 4.267 | 3.7 | 6
| Mean |  | 4.6517 | 6 | 5 | 5 | 5 | 5 | 5.1288 | 4 | 4.4691 | 3.6333 | 6
| Median |  | 4.665 | 6 | 5 | 5 | 5 | 5 | 5.1589 | 4 | 4.3333 | 3.65 | 6
| Std Dev. |  | 0.089981 | 0 | 0 | 0 | 0 | 0 | 0.14405 | 0 | 0.44654 | 0.08165 | 0
| Rel.Std.Dev. |  | 1.934% | 0% | 0% | 0% | 0% | 0% | 2.809% | 0% | 9.992% | 2.247% | 0%
| PDM3 |  | -7.429% | 19.4% | -0.4966% | -0.4966% | -0.4966% | -0.4966% | 2.067% | -20.4% | -11.06% | -27.69% | 19.4%
| Z-Score (Absolute) |  | 0.67 | 1.94 | 0 | 0 | 0 | 0 | 0.25 | 1.94 | 1.03 | 2.65 | 1.94
| Analytical results for W in OREAS 353c (Indicative Value 1.94 ppm) | [15 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 0.62 | 2 | < 10 | < 10 | < 10 | < 10 | 33.45 | < 1 | 2 | 2 | 1.7182 | 0.2 | 1.95 |  | 1
| 1 | 2 | 0.63 | 2.1 | < 10 | < 10 | < 10 | < 10 | 36 | < 1 | 2 | 2 | 1.8606 | 0.2 | 1.82 |  | 8
| 1 | 3 | 0.59 | 2 | < 10 | < 10 | < 10 | < 10 | 35.52 | < 1 | 2 | 2 | 1.7207 | 0.2 | 1.78 |  | 16
| 1 | 4 | 0.66 | 2 | < 10 | < 10 | < 10 | < 10 | 38.45 | < 1 | 2 | 2 | 1.9137 | 0.2 | 1.96 |  | 1.9422
| 1 | 5 | 0.61 | 2.1 | < 10 | < 10 | < 10 | < 10 | 32.72 | < 1 | 2 | 2 | 1.822 | 0.2 | 1.82 |  | 14
| 1 | 6 | 0.61 | 1.9 | < 10 | < 10 | < 10 | < 10 | 32.7 | < 1 | 2 | 2 | 1.6717 | 0.2 | 2.13
| Mean |  | 0.62 | 2.0167 |  |  |  |  | 34.8067 |  | 2 | 2 | 1.7845 | 0.2 | 1.91
| Median |  | 0.615 | 2 |  |  |  |  | 34.485 |  | 2 | 2 | 1.7714 | 0.2 | 1.885
| Std Dev. |  | 0.023664 | 0.075277 |  |  |  |  | 2.2733 |  | 0 | 0 | 0.094933 | 3.0405e-17 | 0.13084
| Rel.Std.Dev. |  | 3.817% | 3.733% |  |  |  |  | 6.531% |  | 0% | 0% | 5.32% | 1.52e-14% | 6.85%
| PDM3 |  | -68.08% | 3.832% |  |  |  |  | 1692% |  | 2.974% | 2.974% | -8.122% | -89.7% | -1.659%
| Z-Score (Absolute) |  | 0.62 | 0.01 | 1.35 | 1.35 | 1.35 | 1.35 | 14.75 | 0.67 | 0 | 0 | 0.1 | 0.81 | 0.04
| Analytical results for Y in OREAS 353c (Indicative Value 5.81 ppm) | [15 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 6.73 | 6.9 | 4.496 | 7.3 | 3.9 | 6.7 | 6.4896 | 6.7 | 4.39 | [5 empty] | 1
| 1 | 2 | 6.65 | 6.8 | 4.6162 | 6.8 | 4.4 | 4.8 | 6.1665 | 6.5 | 4.43 | [5 empty] | 9
| 1 | 3 | 6.47 | 6.9 | 4.3319 | 7.3 | 5.1 | 5.2 | 6.1583 | 7 | 4.45 | [5 empty] | 16
| 1 | 4 | 6.75 | 7.1 | 4.4663 | 8.1 | 4 | 5.1 | 6.0949 | 6.4 | 4.33 | [5 empty] | 5.8109
| 1 | 5 | 6.65 | 7.1 | 4.3543 | 6.7 | 3.8 | 4.8 | 6.4282 | 7 | 4.58 | [5 empty] | 15
| 1 | 6 | 6.6 | 6.7 | 4.5113 | 8.3 | 4 | 4.9 | 6.0577 | 6.6 | 4.69
| Mean |  | 6.6417 | 6.9167 | 4.4627 | 7.4167 | 4.2 | 5.25 | 6.2325 | 6.7 | 4.4783
| Median |  | 6.65 | 6.9 | 4.4812 | 7.3 | 4 | 5 | 6.1624 | 6.65 | 4.44
| Std Dev. |  | 0.10088 | 0.16021 | 0.10575 | 0.65853 | 0.4858 | 0.7287 | 0.18098 | 0.25298 | 0.13273
| Rel.Std.Dev. |  | 1.519% | 2.316% | 2.37% | 8.879% | 11.57% | 13.88% | 2.904% | 3.776% | 2.964%
| PDM3 |  | 14.3% | 19.03% | -23.2% | 27.63% | -27.72% | -9.653% | 7.255% | 15.3% | -22.93%
| Z-Score (Absolute) |  | 0.28 | 0.47 | 1.21 | 0.81 | 1.39 | 0.67 | 0 | 0.32 | 1.2
| Analytical results for Yb in OREAS 353c (Indicative Value 0.28 ppm) | [15 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.08 | 1.09 | 1.10 | 1.11 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 0.13 | 0.25372 | 0.4 | 0.2 | 0.3 | [9 empty] | 1
| 1 | 2 | 0.13 | 0.25599 | 0.4 | 0.3 | 0.4 | [9 empty] | 10
| 1 | 3 | 0.13 | 0.24478 | 0.4 | 0.4 | 0.4 | [9 empty] | 16
| 1 | 4 | 0.13 | 0.26776 | 0.4 | 0.3 | 0.2 | [9 empty] | 0.2838
| 1 | 5 | 0.13 | 0.25409 | 0.4 | 0.3 | 0.3 | [9 empty] | 16
| 1 | 6 | 0.14 | 0.24774 | 0.5 | 0.3 | 0.3
| Mean |  | 0.13167 | 0.25401 | 0.41667 | 0.3 | 0.31667
| Median |  | 0.13 | 0.2539 | 0.4 | 0.3 | 0.3
| Std Dev. |  | 0.0040825 | 0.0079709 | 0.040825 | 0.063246 | 0.075277
| Rel.Std.Dev. |  | 3.101% | 3.138% | 9.798% | 21.08% | 23.77%
| PDM3 |  | -53.61% | -10.5% | 46.82% | 5.707% | 11.58%
| Z-Score (Absolute) |  | 2.47 | 0.67 | 1.71 | 0 | 0.24
| Analytical results for Zn in OREAS 353c (Certified Value 2.79 wt.%) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | 4A*AAS | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.89 | 2.61 | 2.75 | 2.85 | 2.86 | 2.79 | 2.6686 | 2.5849 | 2.8709 | 2.873 | 2.9245 | 2.66 | 2.882 | 2.7944 | 1
| 1 | 2 | 2.88 | 2.67 | 2.72 | 2.85 | 2.86 | 2.78 | 2.6543 | 2.5897 | 2.8536 | 2.856 | 2.9834 | 2.58 | 2.7842 | 2.8296 | #N/A
| 1 | 3 | 2.91 | 2.64 | 2.78 | 2.84 | 2.86 | 2.82 | 2.6698 | 2.5664 | 2.8629 | 2.7834 | 2.9454 | 2.74 | 2.7393 | 2.8027 | 16
| 1 | 4 | 2.87 | 2.58 | 2.74 | 2.86 | 2.88 | 2.8 | 2.634 | 2.5331 | 2.8944 | 2.8258 | 2.9727 | 2.52 | 2.8594 | 2.8157 | 2.7934
| 1 | 5 | 2.89 | 2.62 | 2.75 | 2.82 | 2.87 | 2.79 | 2.6344 | 2.5772 | 2.8163 | 2.8497 | 2.9259 | 2.62 | 2.8461 | 2.7798 | 44
| 1 | 6 | 2.9 | 2.61 | 2.74 | 2.77 | 2.87 | 2.79 | 2.6145 | 2.5968 | 2.872 | 2.8437 | 2.9212 | 2.7 | 2.87 | 2.8043
| Mean |  | 2.89 | 2.6217 | 2.7467 | 2.8317 | 2.8667 | 2.795 | 2.6459 | 2.5747 | 2.8617 | 2.8386 | 2.9455 | 2.6367 | 2.8302 | 2.8044
| Median |  | 2.89 | 2.615 | 2.745 | 2.845 | 2.865 | 2.79 | 2.6443 | 2.5811 | 2.8669 | 2.8467 | 2.9357 | 2.64 | 2.8527 | 2.8035
| Std Dev. |  | 0.014142 | 0.030605 | 0.019664 | 0.033116 | 0.008165 | 0.013784 | 0.021982 | 0.022905 | 0.026032 | 0.031133 | 0.026807 | 0.080416 | 0.056128 | 0.017151
| Rel.Std.Dev. |  | 0.4893% | 1.167% | 0.7159% | 1.169% | 0.2848% | 0.4932% | 0.8308% | 0.8896% | 0.9097% | 1.097% | 0.9101% | 3.05% | 1.983% | 0.6116%
| PDM3 |  | 3.457% | -6.149% | -1.674% | 1.369% | 2.621% | 0.05594% | -5.28% | -7.831% | 2.443% | 1.617% | 5.445% | -5.612% | 1.315% | 0.393%
| Z-Score (Absolute) |  | 0.82 | 2.2 | 0.79 | 0.16 | 0.55 | 0.25 | 1.93 | 2.73 | 0.5 | 0.24 | 1.44 | 2.03 | 0.14 | 0.14
| Analytical results for Zr in OREAS 353c (Indicative Value 2.53 ppm) | [15 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 2.77 | 2.8 | 3.2 | 2 | 3 | 1.8796 | 2.2 | 2.7 | [6 empty] | 1
| 1 | 2 | 2.81 | 2.7 | 3.8 | 2 | 2 | 1.9467 | 2.3 | 3 | [6 empty] | 11
| 1 | 3 | 2.93 | 2.8 | 4.2 | 2 | 2 | 1.9839 | 2.4 | 2.3 | [6 empty] | 16
| 1 | 4 | 2.8 | 2.8 | 3.4 | 2 | 2 | 1.8215 | 2.1 | 2.3 | [6 empty] | 2.5283
| 1 | 5 | 2.76 | 2.7 | 3.6 | 2 | 2 | 1.9998 | 2.4 | 2.6 | [6 empty] | 17
| 1 | 6 | 2.89 | 2.6 | 4.2 | 2 | 2 | 1.8663 | 2.4 | 2.4
| Mean |  | 2.8267 | 2.7333 | 3.7333 | 2 | 2.1667 | 1.9163 | 2.3 | 2.55
| Median |  | 2.805 | 2.75 | 3.7 | 2 | 2 | 1.9131 | 2.35 | 2.5
| Std Dev. |  | 0.068313 | 0.08165 | 0.41312 | 0 | 0.40825 | 0.071155 | 0.12649 | 0.27386
| Rel.Std.Dev. |  | 2.417% | 2.987% | 11.07% | 0% | 18.84% | 3.713% | 5.5% | 10.74%
| PDM3 |  | 11.8% | 8.11% | 47.66% | -20.9% | -14.3% | -24.21% | -9.029% | 0.8588%
| Z-Score (Absolute) |  | 0.76 | 0.59 | 2.49 | 0.81 | 0.49 | 0.97 | 0.24 | 0.24
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C707: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C725: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C743: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C761: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C833: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C851: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C869: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C887: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C905: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C923: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C959: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C977: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C995: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1013: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1031: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 353c (Indicative Value 0.275 wt.%) | [15 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.08 | 1.10 | 1.11 | 1.13 | [10 empty] | wt.%
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | [10 empty] | 3
3
| 1 | 1 | 0.30812 | 0.2846 | 0.2432 | 0.261 | [10 empty] | 1
| 1 | 2 | 0.30798 | 0.277 | 0.2504 | 0.258 | [10 empty] | 13
| 1 | 3 | 0.30869 | 0.2855 | 0.2396 | 0.273 | [10 empty] | 16
| 1 | 4 | 0.30943 | 0.2761 | 0.2417 | 0.265 | [10 empty] | 0.27483
| 1 | 5 | 0.30831 | 0.2817 | 0.2456 | 0.264 | [10 empty] | 19
| 1 | 6 | 0.30915 | 0.2862 | 0.2487 | 0.263
| Mean |  | 0.30861 | 0.28185 | 0.24487 | 0.264
| Median |  | 0.3085 | 0.28315 | 0.2444 | 0.2635
| Std Dev. |  | 0.00058234 | 0.0043912 | 0.0041568 | 0.0050596
| Rel.Std.Dev. |  | 0.1887% | 1.558% | 1.698% | 1.917%
| PDM3 |  | 12.29% | 2.553% | -10.9% | -3.941%
| Z-Score (Absolute) |  | 1.3 | 0.33 | 1.02 | 0.33
| Analytical results for Al in OREAS 353c (Certified Value 0.1 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 3
3
| 1 | 1 | 0.11 | 0.1 | 0.101 | 0.08 | 0.148 | 0.09 | 0.111 | 0.13489 | 0.084 | 0.1 | 0.11 | 0.097732 | 0.1 | 0.097 | 1
| 1 | 2 | 0.1 | 0.12 | 0.116 | 0.09 | 0.111 | 0.101 | 0.101 | 0.13206 | 0.098 | 0.11 | 0.08 | 0.088849 | 0.1 | 0.0958 | #N/A
| 1 | 3 | 0.1 | 0.11 | 0.111 | 0.09 | 0.122 | 0.09 | 0.101 | 0.13911 | 0.105 | 0.1 | 0.07 | 0.088466 | 0.11 | 0.1316 | 16
| 1 | 4 | 0.11 | 0.11 | 0.106 | 0.1 | 0.122 | 0.09 | 0.101 | 0.13466 | 0.102 | 0.11 | 0.17 | 0.11133 | 0.1 | 0.0933 | 0.099907
| 1 | 5 | 0.1 | 0.1 | 0.122 | 0.09 | 0.132 | 0.09 | 0.111 | 0.131 | 0.094 | 0.11 | 0.09 | 0.11218 | 0.1 | 0.0976 | 46
| 1 | 6 | 0.1 | 0.1 | 0.122 | 0.11 | 0.138 | 0.101 | 0.101 | 0.13308 | 0.098 | 0.11 | 0.07 | 0.10186 | 0.1 | 0.0954
| Mean |  | 0.10333 | 0.10667 | 0.113 | 0.093333 | 0.12883 | 0.093667 | 0.10433 | 0.13413 | 0.096833 | 0.10667 | 0.098333 | 0.10007 | 0.10167 | 0.10178
| Median |  | 0.1 | 0.105 | 0.1135 | 0.09 | 0.127 | 0.09 | 0.101 | 0.13387 | 0.098 | 0.11 | 0.085 | 0.099797 | 0.1 | 0.0964
| Std Dev. |  | 0.005164 | 0.008165 | 0.008579 | 0.010328 | 0.013212 | 0.0056804 | 0.005164 | 0.0028598 | 0.0073326 | 0.005164 | 0.038166 | 0.010419 | 0.0040825 | 0.014683
| Rel.Std.Dev. |  | 4.997% | 7.655% | 7.592% | 11.07% | 10.26% | 6.064% | 4.949% | 2.132% | 7.572% | 4.841% | 38.81% | 10.41% | 4.016% | 14.43%
| PDM3 |  | 3.43% | 6.766% | 13.11% | -6.579% | 28.95% | -6.246% | 4.431% | 34.26% | -3.076% | 6.766% | -1.575% | 0.1644% | 1.762% | 1.879%
| Z-Score (Absolute) |  | 0.13 | 0.66 | 1.67 | 1.49 | 4.24 | 1.47 | 0.25 | 5.11 | 0.93 | 0.66 | 0.68 | 0.4 | 0.14 | 0.13
| Analytical results for As in OREAS 353c (Certified Value 292 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 316 | 290 | 400 | 400 | 400 | 300 | 300 | 313 | 313 | 281 | 281.6925 | 272 |  |  | 1
| 1 | 2 | 314 | 290 | 400 | 300 | 400 | 400 | 300 | 310 | 302 | 276 | 295.8912 | 269 |  |  | #N/A
| 1 | 3 | 318 | 280 | 400 | 400 | 400 | 300 | 300 | 279 | 308 | 276 | 305.653 | 277 |  |  | 16
| 1 | 4 | 320 | 280 | 400 | 300 | 400 | 400 | 300 | 274 | 296 | 279 | 284.8048 | 262 |  |  | 291.6764
| 1 | 5 | 325 | 280 | 400 | 400 | 400 | 300 | 300 | 276 | 308 | 281 | 308.592 | 271 |  |  | 47
| 1 | 6 | 314 | 290 | 400 | 400 | 400 | 400 | 300 | 282 | 310 | 271 | 297.7749 | 273
| Mean |  | 317.8333 | 285 | 400 | 366.6667 | 400 | 350 | 300 | 289 | 306.1667 | 277.3333 | 295.7348 | 270.6667
| Median |  | 317 | 285 | 400 | 400 | 400 | 350 | 300 | 280.5 | 308 | 277.5 | 296.833 | 271.5
| Std Dev. |  | 4.2151 | 5.4772 | 0 | 51.6398 | 0 | 54.7723 | 0 | 17.6635 | 6.1455 | 3.8297 | 10.8122 | 5.0067
| Rel.Std.Dev. |  | 1.326% | 1.922% | 0% | 14.08% | 0% | 15.65% | 0% | 6.112% | 2.007% | 1.381% | 3.656% | 1.85%
| PDM3 |  | 8.968% | -2.289% | 37.14% | 25.71% | 37.14% | 20% | 2.854% | -0.9176% | 4.968% | -4.917% | 1.391% | -7.203%
| Z-Score (Absolute) |  | 1.67 | 0.23 | NA | NA | NA | NA | NA | 0 | 0.99 | 0.67 | 0.39 | 1.06
| Analytical results for B in OREAS 353c (Indicative Value < 50 ppm) | [15 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.12 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [9 empty] | 0
0
| 1 | 1 | 27 | < 50 | < 50 | < 20 | < 30 | [9 empty] | 1
| 1 | 2 | 27 | < 50 | < 50 | < 20 | < 30 | [9 empty] | 14
| 1 | 3 | 27 | < 50 | < 50 | < 20 | < 30 | [9 empty] | 16
| 1 | 4 | 28 | < 50 | < 50 | < 20 | < 30 | [9 empty] | < 50
| 1 | 5 | 27 | < 50 | < 50 | < 20 | < 30 | [9 empty] | 20
| 1 | 6 | 29 | < 50 | < 50 | < 20 | < 30
| Mean |  | 27.5
| Median |  | 27
| Std Dev. |  | 0.83666
| Rel.Std.Dev. |  | 3.042%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 4.05 | 2.7
| Analytical results for Ba in OREAS 353c (Certified Value 38.6 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [6 empty] | 1
1
| 1 | 1 | 31 | 36 | 23 | 44 | 37 | 40.2139 | 42 | 63 | [6 empty] | 1
| 1 | 2 | 33 | 39 | 14 | 43 | 33 | 40.7059 | 42 | 46 | [6 empty] | #N/A
| 1 | 3 | 29 | 38 | 17 | 42 | 37 | 41.8599 | 44 | 61 | [6 empty] | 16
| 1 | 4 | 31 | 37 | 14 | 41 | 42 | 41.9881 | 42 | 44 | [6 empty] | 38.6037
| 1 | 5 | 31 | 38 | 65 | 45 | 36 | 42.2886 | 41 | 47 | [6 empty] | 48
| 1 | 6 | 30 | 37 | 31 | 45 | 35 | 40.6775 | 41 | 65
| Mean |  | 30.8333 | 37.5 | 27.3333 | 43.3333 | 36.6667 | 41.289 | 42 | 54.3333
| Median |  | 31 | 37.5 | 20 | 43.5 | 36.5 | 41.2829 | 42 | 54
| Std Dev. |  | 1.3292 | 1.0488 | 19.5619 | 1.633 | 3.0111 | 0.85833 | 1.0954 | 9.6264
| Rel.Std.Dev. |  | 4.311% | 2.797% | 71.57% | 3.768% | 8.212% | 2.079% | 2.608% | 17.72%
| PDM3 |  | -20.13% | -2.859% | -29.2% | 12.25% | -5.018% | 6.956% | 8.798% | 40.75%
| Z-Score (Absolute) |  | 1.73 | 0.38 | 2.44 | 0.8 | 0.55 | 0.38 | 0.53 | 3.02
| Analytical results for Be in OREAS 353c (Indicative Value < 1 ppm) | [15 empty] | Indicative
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.12 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | < 0.5 | < 1 | 0.1568 | < 5 | < 1 | < 20 | 2.3 | [7 empty] | 1
| 1 | 2 | < 0.5 | < 1 | 0.1539 | < 5 | < 1 | < 20 | 2.2 | [7 empty] | 15
| 1 | 3 | < 0.5 | < 1 | 0.1434 | < 5 | < 1 | < 20 | 2.9 | [7 empty] | 16
| 1 | 4 | < 0.5 | < 1 | 0.1431 | < 5 | < 1 | < 20 | 2.1 | [7 empty] | < 1
| 1 | 5 | < 0.5 | < 1 | 0.1589 | < 5 | < 1 | < 20 | 1.7 | [7 empty] | 21
| 1 | 6 | < 0.5 | < 1 | 0.1505 | < 5 | < 1 | < 20 | 1.4
| Mean |  |  |  | 0.1511 |  |  |  | 2.1
| Median |  |  |  | 0.1522 |  |  |  | 2.15
| Std Dev. |  |  |  | 0.0067055 |  |  |  | 0.51769
| Rel.Std.Dev. |  |  |  | 4.438% |  |  |  | 24.65%
PDM3
| Z-Score (Absolute) |  | 0.48 | 0 | 0.67 | 3.87 | 0 | 18.36 | 3.09
| Analytical results for Bi in OREAS 353c (Certified Value 111 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 112.07 | 116 | 120 | 112.3 | 106 | 105.5751 | 107 | [7 empty] | 1
| 1 | 2 | 113.42 | 118 | 120.03 | 109.4 | 105.2 | 106.7206 | 107 | [7 empty] | #N/A
| 1 | 3 | 112.51 | 117 | 114.6 | 111.1 | 106.5 | 105.7441 | 107 | [7 empty] | 16
| 1 | 4 | 117.81 | 118 | 115.02 | 105.1 | 107 | 110.2126 | 104 | [7 empty] | 111.3597
| 1 | 5 | 115.79 | 117 | 119.21 | 110.3 | 104.9 | 110.2627 | 105 | [7 empty] | 49
| 1 | 6 | 115.33 | 119 | 118.59 | 113.1 | 107.6 | 106.7129 | 104
| Mean |  | 114.4883 | 117.5 | 117.9083 | 110.2167 | 106.2 | 107.538 | 105.6667
| Median |  | 114.375 | 117.5 | 118.9 | 110.7 | 106.25 | 106.7168 | 106
| Std Dev. |  | 2.2063 | 1.0488 | 2.4626 | 2.8386 | 1.0412 | 2.1446 | 1.5055
| Rel.Std.Dev. |  | 1.927% | 0.8926% | 2.089% | 2.575% | 0.9804% | 1.994% | 1.425%
| PDM3 |  | 2.809% | 5.514% | 5.881% | -1.026% | -4.633% | -3.432% | -5.112%
| Z-Score (Absolute) |  | 0.67 | 1.15 | 1.21 | 0 | 0.63 | 0.42 | 0.72
| Analytical results for Ca in OREAS 353c (Certified Value 0.337 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 3
3
| 1 | 1 | 0.38 | 0.3 | 0.257 | 0.279 | 0.329 | 0.35 | 0.429 | 0.31201 | 0.343 | 0.3 | 0.4 | 0.35754 | 0.38 | 0.5011 | 1
| 1 | 2 | 0.38 | 0.3 | 0.329 | 0.322 | 0.307 | 0.343 | 0.35 | 0.31604 | 0.35 | 0.3 | 0.4 | 0.3457 | 0.36 | 0.4447 | #N/A
| 1 | 3 | 0.37 | 0.3 | 0.307 | 0.272 | 0.343 | 0.35 | 0.329 | 0.3107 | 0.343 | 0.3 | 0.4 | 0.33269 | 0.37 | 0.5747 | 16
| 1 | 4 | 0.37 | 0.3 | 0.257 | 0.393 | 0.336 | 0.357 | 0.293 | 0.31033 | 0.346 | 0.3 | 0.4 | 0.36873 | 0.35 | 0.3534 | 0.33709
| 1 | 5 | 0.36 | 0.3 | 0.329 | 0.35 | 0.336 | 0.35 | 0.329 | 0.31474 | 0.337 | 0.3 | 0.4 | 0.35336 | 0.35 | 0.4678 | 50
| 1 | 6 | 0.37 | 0.3 | 0.322 | 0.322 | 0.322 | 0.343 | 0.343 | 0.31687 | 0.367 | 0.3 | 0.4 | 0.33384 | 0.35 | 0.4482
| Mean |  | 0.37167 | 0.3 | 0.30017 | 0.323 | 0.32883 | 0.34883 | 0.3455 | 0.31345 | 0.34767 | 0.3 | 0.4 | 0.34864 | 0.36 | 0.46498
| Median |  | 0.37 | 0.3 | 0.3145 | 0.322 | 0.3325 | 0.35 | 0.336 | 0.31337 | 0.3445 | 0.3 | 0.4 | 0.34953 | 0.355 | 0.458
| Std Dev. |  | 0.0075277 | 0 | 0.034388 | 0.045095 | 0.012859 | 0.0052694 | 0.045386 | 0.0028074 | 0.010386 | 0 | 6.0809e-17 | 0.014054 | 0.012649 | 0.072793
| Rel.Std.Dev. |  | 2.025% | 0% | 11.46% | 13.96% | 3.911% | 1.511% | 13.14% | 0.8956% | 2.987% | 0% | 1.52e-14% | 4.031% | 3.514% | 15.65%
| PDM3 |  | 10.26% | -11% | -10.95% | -4.181% | -2.45% | 3.483% | 2.494% | -7.014% | 3.137% | -11% | 18.66% | 3.427% | 6.796% | 37.94%
| Z-Score (Absolute) |  | 0.86 | NA | 1.69 | 0.89 | 0.67 | 0.05 | 0.08 | 1.22 | 0 | NA | NA | 0.03 | 0.44 | 4.18
Results from laboratories 2, 10 and 11 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 353c (Certified Value 154 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 0
0
| 1 | 1 | 156.09 | 160 | 158.312 | 170 | 167 | 152 | 145.2055 | 150 | [6 empty] | 1
| 1 | 2 | 153.83 | 150 | 163.328 | 164 | 165 | 146 | 138.5191 | 149 | [6 empty] | #N/A
| 1 | 3 | 152.43 | 150 | 155.24 | 168 | 168 | 146 | 140.94 | 148 | [6 empty] | 16
| 1 | 4 | 153.14 | 160 | 159.272 | 162 | 159 | 142 | 145.7888 | 146 | [6 empty] | 154.345
| 1 | 5 | 155.15 | 160 | 158.768 | 161 | 162 | 150 | 142.8889 | 146 | [6 empty] | 51
| 1 | 6 | 154.62 | 160 | 159.944 | 156 | 166 | 146 | 140.093 | 146
| Mean |  | 154.21 | 156.6667 | 159.144 | 163.5 | 164.5 | 147 | 142.2392 | 147.5
| Median |  | 154.225 | 160 | 159.02 | 163 | 165.5 | 146 | 141.9145 | 147
| Std Dev. |  | 1.3449 | 5.164 | 2.6168 | 5.0498 | 3.3912 | 3.5214 | 2.8967 | 1.7607
| Rel.Std.Dev. |  | 0.8721% | 3.296% | 1.644% | 3.089% | 2.061% | 2.395% | 2.037% | 1.194%
| PDM3 |  | -0.08746% | 1.504% | 3.109% | 5.932% | 6.579% | -4.759% | -7.843% | -4.435%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 0.31 | 0.68 | 0.76 | 0.71 | 1.11 | 0.67
| Analytical results for Ce in OREAS 353c (Indicative Value 16.5 ppm) | [15 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.11 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 1
1
| 1 | 1 | 16.82 | 16.7 | 16.7 | 16.8 | 48 | [9 empty] | 1
| 1 | 2 | 16.56 | 16.4 | 16.1 | 16.7 | 22 | [9 empty] | 16
| 1 | 3 | 16.62 | 16.8 | 16.3 | 16.7 | 66 | [9 empty] | 16
| 1 | 4 | 16.65 | 17.1 | 16.9 | 16.1 | 12 | [9 empty] | 16.5183
| 1 | 5 | 16.54 | 16.3 | 16 | 15.9 | 12 | [9 empty] | 22
| 1 | 6 | 16.35 | 16.6 | 16.4 | 16.4 | 42
| Mean |  | 16.59 | 16.65 | 16.4 | 16.4333 | 33.6667
| Median |  | 16.59 | 16.65 | 16.35 | 16.55 | 32
| Std Dev. |  | 0.15388 | 0.2881 | 0.34641 | 0.36697 | 21.8876
| Rel.Std.Dev. |  | 0.9276% | 1.73% | 2.112% | 2.233% | 65.01%
| PDM3 |  | 0.4339% | 0.7971% | -0.7164% | -0.5146% | 103.8%
| Z-Score (Absolute) |  | 0 | 0.26 | 0.82 | 0.67 | 73.5
| Analytical results for Co in OREAS 353c (Certified Value 18.1 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES/MS |  |  | 1
1
| 1 | 1 | 17.21 | 20 | < 20 | 30 | < 20 | < 20 | 40 | 14.205 | 20 | 20 | 19.5935 | 18 |  |  | 1
| 1 | 2 | 16.65 | 20 | < 20 | 30 | < 20 | < 20 | < 20 | 14.299 | 20 | 20 | 19.9814 | 18 |  |  | #N/A
| 1 | 3 | 17.14 | 20 | < 20 | 20 | < 20 | < 20 | < 20 | 14.26 | 18 | 20 | 19.1598 | 18 |  |  | 16
| 1 | 4 | 17.06 | 20 | < 20 | < 20 | < 20 | < 20 | 20 | 14.23 | 17 | 19 | 20.3807 | 18 |  |  | 18.0871
| 1 | 5 | 17.14 | 20 | < 20 | 20 | < 20 | < 20 | < 20 | 14.352 | 17 | 20 | 19.4172 | 18 |  |  | 52
| 1 | 6 | 16.48 | 20 | < 20 | < 20 | < 20 | 140 | < 20 | 14.391 | 16 | 20 | 18.7093 | 18
| Mean |  | 16.9467 | 20 |  | 25 |  | 140 | 30 | 14.2895 | 18 | 19.8333 | 19.5403 | 18
| Median |  | 17.1 | 20 |  | 25 |  | 140 | 30 | 14.2795 | 17.5 | 20 | 19.5054 | 18
| Std Dev. |  | 0.30421 | 0 |  | 5.7735 |  |  | 14.1421 | 0.071902 | 1.6733 | 0.40825 | 0.59232 | 0
| Rel.Std.Dev. |  | 1.795% | 0% |  | 23.09% |  |  | 47.14% | 0.5032% | 9.296% | 2.058% | 3.031% | 0%
| PDM3 |  | -6.305% | 10.58% |  | 38.22% |  | 674% | 65.86% | -21% | -0.4817% | 9.654% | 8.034% | -0.4817%
| Z-Score (Absolute) |  | 0.39 | 0.74 | 2.94 | NA | 2.94 | NA | NA | 1.36 | 0 | 0.67 | 0.57 | 0
| Analytical results for Cr in OREAS 353c (Certified Value 117 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS |  |  | 0
0
| 1 | 1 | 108 | 130 | 140 | 70 | 210 | 210 | 70 | 121 | 140 | 134.426 | 130 | 128 |  |  | 1
| 1 | 2 | 113 | 130 | 140 | 70 | 210 | 210 | 70 | 125 | 141 | 122.4156 | 130 | 129 |  |  | #N/A
| 1 | 3 | 114 | 130 | 70 | 70 | 70 | 210 | 70 | 121 | 139 | 124.3667 | 140 | 130 |  |  | 16
| 1 | 4 | 108 | 140 | 140 | 210 | 210 | 210 | 70 | 111 | 137 | 135.0947 | 130 | 123 |  |  | 117.446
| 1 | 5 | 110 | 140 | 140 | 70 | 210 | 210 | 70 | 129 | 139 | 131.5619 | 130 | 128 |  |  | 53
| 1 | 6 | 111 | 130 | 140 | 70 | 210 | 210 | 70 | 120 | 149 | 128.6803 | 130 | 119
| Mean |  | 110.6667 | 133.3333 | 128.3333 | 93.3333 | 186.6667 | 210 | 70 | 121.1667 | 140.8333 | 129.4242 | 131.6667 | 126.1667
| Median |  | 110.5 | 130 | 140 | 70 | 210 | 210 | 70 | 121 | 139.5 | 130.1211 | 130 | 128
| Std Dev. |  | 2.5033 | 5.164 | 28.5774 | 57.1548 | 57.1548 | 0 | 0 | 6.0139 | 4.2151 | 5.2344 | 4.0825 | 4.2622
| Rel.Std.Dev. |  | 2.262% | 3.873% | 22.27% | 61.24% | 30.62% | 0% | 0% | 4.963% | 2.993% | 4.044% | 3.101% | 3.378%
| PDM3 |  | -5.772% | 13.53% | 9.27% | -20.53% | 58.94% | 78.81% | -40.4% | 3.168% | 19.91% | 10.2% | 12.11% | 7.425%
| Z-Score (Absolute) |  | 0.68 | NA | 0.02 | 1.55 | 2.53 | 3.55 | 2.57 | 0.21 | 0.67 | 0.16 | NA | 0.02
Results from laboratories 2 and 13 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 353c (Indicative Value 0.21 ppm) | [15 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [7 empty] | 2
2
| 1 | 1 | < 1 | 0.23 | 0.1596 | 0.2 | 0.3 | 0.18969 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 1 | 0.25 | 0.1448 | 0.2 | 0.2 | 0.1727 | < 0.1 | [7 empty] | 17
| 1 | 3 | < 1 | 0.21 | 0.1474 | 0.2 | 0.2 | 0.17609 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 1 | 0.23 | 0.1517 | 0.2 | 0.3 | 0.16582 | < 0.1 | [7 empty] | 0.20519
| 1 | 5 | < 1 | 0.21 | 0.1551 | 0.4 | 0.1 | 0.19186 | < 0.1 | [7 empty] | 23
| 1 | 6 | < 1 | 0.22 | 0.138 | 0.2 | 0.2 | 0.17831 | < 0.1
| Mean |  |  | 0.225 | 0.14943 | 0.23333 | 0.21667 | 0.17908
| Median |  |  | 0.225 | 0.14955 | 0.2 | 0.2 | 0.1772
| Std Dev. |  |  | 0.015166 | 0.0077063 | 0.08165 | 0.075277 | 0.010019
| Rel.Std.Dev. |  |  | 6.74% | 5.157% | 34.99% | 34.74% | 5.595%
| PDM3 |  |  | 9.657% | -27.17% | 13.72% | 5.595% | -12.72%
| Z-Score (Absolute) |  | 5.08 | 0.15 | 1.21 | 0.3 | 0 | 0.67 | 2.99
| Analytical results for Cu in OREAS 353c (Certified Value 0.348 wt.%) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 |  | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS |  | 3
3
| 1 | 1 | 0.3729 | 0.348 | 0.334 | 0.35 | 0.357 | 0.347 | 0.33 | 0.36684 | 0.364 | 0.3482 | 0.35943 | 0.344 | 0.3525 |  | 1
| 1 | 2 | 0.3677 | 0.344 | 0.337 | 0.345 | 0.335 | 0.347 | 0.335 | 0.3639 | 0.3688 | 0.3447 | 0.35385 | 0.336 | 0.3582 |  | #N/A
| 1 | 3 | 0.3672 | 0.344 | 0.331 | 0.347 | 0.358 | 0.346 | 0.334 | 0.36137 | 0.3646 | 0.3497 | 0.35626 | 0.349 | 0.3542 |  | 16
| 1 | 4 | 0.3755 | 0.341 | 0.33 | 0.328 | 0.356 | 0.343 | 0.324 | 0.36824 | 0.3664 | 0.3365 | 0.35709 | 0.338 | 0.3469 |  | 0.34822
| 1 | 5 | 0.3726 | 0.349 | 0.336 | 0.35 | 0.352 | 0.347 | 0.331 | 0.3643 | 0.3638 | 0.3459 | 0.35917 | 0.339 | 0.3606 |  | 54
| 1 | 6 | 0.3746 | 0.344 | 0.335 | 0.346 | 0.358 | 0.346 | 0.335 | 0.36809 | 0.3639 | 0.3509 | 0.35491 | 0.337 | 0.3353
| Mean |  | 0.37175 | 0.345 | 0.33383 | 0.34433 | 0.35267 | 0.346 | 0.3315 | 0.36545 | 0.36525 | 0.34598 | 0.35678 | 0.3405 | 0.35128
| Median |  | 0.37275 | 0.344 | 0.3345 | 0.3465 | 0.3565 | 0.3465 | 0.3325 | 0.36557 | 0.3643 | 0.34705 | 0.35667 | 0.3385 | 0.35335
| Std Dev. |  | 0.0035024 | 0.0029665 | 0.0027869 | 0.0082624 | 0.0089368 | 0.0015492 | 0.0042308 | 0.0027234 | 0.0019917 | 0.0051863 | 0.0022463 | 0.00501 | 0.0091576
| Rel.Std.Dev. |  | 0.9421% | 0.8598% | 0.8348% | 2.4% | 2.534% | 0.4477% | 1.276% | 0.7452% | 0.5453% | 1.499% | 0.6296% | 1.471% | 2.607%
| PDM3 |  | 6.759% | -0.9235% | -4.13% | -1.115% | 1.278% | -0.6363% | -4.8% | 4.951% | 4.892% | -0.6411% | 2.461% | -2.216% | 0.8809%
| Z-Score (Absolute) |  | 2.6 | 0.1 | 1.23 | 0.17 | 0.67 | 0 | 1.47 | 1.97 | 1.95 | 0 | 1.09 | 0.56 | 0.53
| Analytical results for Dy in OREAS 353c (Indicative Value 1.33 ppm) | [15 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.11 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.94 | 1.5 | 1.2 | 1.42 | 2 | [9 empty] | 1
| 1 | 2 | 0.85 | 1.5 | 1.3 | 1.32 | 2 | [9 empty] | 3
| 1 | 3 | 0.92 | 1.5 | 1.3 | 1.35 | 2 | [9 empty] | 16
| 1 | 4 | 0.9 | 1.4 | 1.2 | 1.36 | < 2 | [9 empty] | 1.325
| 1 | 5 | 0.89 | 1.4 | 1.3 | 1.28 | < 2 | [9 empty] | 24
| 1 | 6 | 0.92 | 1.5 | 1.2 | 1.3 | 2
| Mean |  | 0.90333 | 1.4667 | 1.25 | 1.3383 | 2
| Median |  | 0.91 | 1.5 | 1.25 | 1.335 | 2
| Std Dev. |  | 0.031411 | 0.05164 | 0.054772 | 0.049967 | 0
| Rel.Std.Dev. |  | 3.477% | 3.521% | 4.382% | 3.733% | 0%
| PDM3 |  | -31.82% | 10.69% | -5.66% | 1.006% | 50.94%
| Z-Score (Absolute) |  | 2.29 | 0.67 | 0.46 | 0 | 1.73
| Analytical results for Er in OREAS 353c (Certified Value 0.62 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 0.55 | 0.7 | 0.5257 | 0.7 | 0.69 | < 2 | [8 empty] | 1
| 1 | 2 | 0.52 | 0.7 | 0.5279 | 0.6 | 0.68 | < 2 | [8 empty] | 4
| 1 | 3 | 0.52 | 0.7 | 0.5189 | 0.7 | 0.63 | < 2 | [8 empty] | 16
| 1 | 4 | 0.55 | 0.7 | 0.5191 | 0.7 | 0.73 | < 2 | [8 empty] | 0.62335
| 1 | 5 | 0.53 | 0.7 | 0.538 | 0.6 | 0.73 | < 2 | [8 empty] | 55
| 1 | 6 | 0.54 | 0.8 | 0.5209 | 0.6 | 0.68 | < 2
| Mean |  | 0.535 | 0.71667 | 0.52508 | 0.65 | 0.69
| Median |  | 0.535 | 0.7 | 0.5233 | 0.65 | 0.685
| Std Dev. |  | 0.013784 | 0.040825 | 0.0073033 | 0.054772 | 0.037417
| Rel.Std.Dev. |  | 2.576% | 5.696% | 1.391% | 8.427% | 5.423%
| PDM3 |  | -14.17% | 14.97% | -15.76% | 4.275% | 10.69%
| Z-Score (Absolute) |  | 1 | 0.35 | 1.08 | 0.15 | 0.15 | 2.45
| Analytical results for Eu in OREAS 353c (Indicative Value 1.9 ppm) | [15 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 1.26 | 2.5 | 1.8766 | 2 | 2.13 | < 2 | [8 empty] | 1
| 1 | 2 | 1.27 | 2.5 | 1.8754 | 2.2 | 2.13 | < 2 | [8 empty] | 5
| 1 | 3 | 1.29 | 2.5 | 1.949 | 2 | 2.19 | 3 | [8 empty] | 16
| 1 | 4 | 1.3 | 2.5 | 1.8297 | 2.2 | 2.23 | < 2 | [8 empty] | 1.8999
| 1 | 5 | 1.27 | 2.4 | 1.9032 | 2 | 2.14 | 2 | [8 empty] | 25
| 1 | 6 | 1.28 | 2.4 | 1.9128 | 2.2 | 2.16 | < 2
| Mean |  | 1.2783 | 2.4667 | 1.8911 | 2.1 | 2.1633 | 2.5
| Median |  | 1.275 | 2.5 | 1.8899 | 2.1 | 2.15 | 2.5
| Std Dev. |  | 0.01472 | 0.05164 | 0.040477 | 0.10954 | 0.039833 | 0.70711
| Rel.Std.Dev. |  | 1.151% | 2.094% | 2.14% | 5.216% | 1.841% | 28.28%
| PDM3 |  | -32.72% | 29.83% | -0.4635% | 10.53% | 13.87% | 31.59%
| Z-Score (Absolute) |  | 1.51 | 0.99 | 0.22 | 0.22 | 0.35 | 1.05
| Analytical results for Fe in OREAS 353c (Certified Value 2.95 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS | 2
3
| 1 | 1 | 2.97 | 3 | 2.917 | 2.92 | 2.917 | 2.952 | 2.952 | 3.06 | 2.938 | 2.99 | 2.57 | 3.1274 | 2.98 | 2.8123 | 1
| 1 | 2 | 2.99 | 2.98 | 3.001 | 2.89 | 2.84 | 2.973 | 2.889 | 3.0831 | 3.044 | 2.96 | 2.49 | 2.9774 | 2.92 | 2.8755 | #N/A
| 1 | 3 | 2.94 | 2.96 | 2.819 | 2.94 | 2.882 | 2.959 | 2.931 | 3.0593 | 3.001 | 3 | 2.47 | 3.1172 | 3.06 | 2.8664 | 16
| 1 | 4 | 2.95 | 2.96 | 2.861 | 2.92 | 2.973 | 2.917 | 2.798 | 3.0319 | 2.949 | 2.9 | 2.55 | 2.9783 | 2.95 | 2.7541 | 2.9494
| 1 | 5 | 2.96 | 3.03 | 2.917 | 2.93 | 2.903 | 2.917 | 2.861 | 3.0586 | 2.892 | 2.98 | 2.52 | 3.1483 | 2.96 | 2.8966 | 56
| 1 | 6 | 2.98 | 2.97 | 2.91 | 2.93 | 2.959 | 2.938 | 2.882 | 3.0246 | 2.808 | 3.02 | 2.49 | 3.0358 | 2.97 | 2.7446
| Mean |  | 2.965 | 2.9833 | 2.9042 | 2.9217 | 2.9123 | 2.9427 | 2.8855 | 3.0529 | 2.9387 | 2.975 | 2.515 | 3.0641 | 2.9733 | 2.8249
| Median |  | 2.965 | 2.975 | 2.9135 | 2.925 | 2.91 | 2.945 | 2.8855 | 3.059 | 2.9435 | 2.985 | 2.505 | 3.0765 | 2.965 | 2.8394
| Std Dev. |  | 0.018708 | 0.027325 | 0.061411 | 0.017224 | 0.049233 | 0.022862 | 0.054357 | 0.021326 | 0.082873 | 0.041833 | 0.038859 | 0.076941 | 0.047188 | 0.064874
| Rel.Std.Dev. |  | 0.631% | 0.9159% | 2.115% | 0.5895% | 1.69% | 0.7769% | 1.884% | 0.6985% | 2.82% | 1.406% | 1.545% | 2.511% | 1.587% | 2.296%
| PDM3 |  | 0.5284% | 1.15% | -1.534% | -0.9408% | -1.257% | -0.2288% | -2.167% | 3.51% | -0.3644% | 0.8675% | -14.73% | 3.888% | 0.811% | -4.221%
| Z-Score (Absolute) |  | 0.46 | 0.81 | 0.69 | 0.36 | 0.54 | 0.03 | 1.05 | 2.13 | 0.03 | 0.65 | 8.07 | 2.34 | 0.62 | 2.19
| Analytical results for Ga in OREAS 353c (Certified Value 1.02 ppm) | [15 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.10 | 1.11 | 1.12 | 1.13 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 1.01 | 1 |  | 3 | 1.0054 | 1 | [8 empty] | 1
| 1 | 2 | 1.01 | 1 | 1 | 3 | 1.0283 | 1 | [8 empty] | #N/A
| 1 | 3 | 1.04 | 1 |  | 3 | 1.0907 | 0.9 | [8 empty] | 16
| 1 | 4 | 1.11 | 1 |  | 4 | 1.0547 | 0.9 | [8 empty] | 1.0156
| 1 | 5 | 1.12 | 1 | 1 | 3 | 1.0675 | 1 | [8 empty] | 57
| 1 | 6 | 1 | 1 | 1 | 2 | 1.0319 | 1.1
| Mean |  | 1.0483 | 1 | 1 | 3 | 1.0464 | 0.98333
| Median |  | 1.025 | 1 | 1 | 3 | 1.0433 | 1
| Std Dev. |  | 0.053448 | 0 | 0 | 0.63246 | 0.030657 | 0.075277
| Rel.Std.Dev. |  | 5.098% | 0% | 0% | 21.08% | 2.93% | 7.655%
| PDM3 |  | 3.221% | -1.538% | -1.538% | 195.4% | 3.032% | -3.179%
| Z-Score (Absolute) |  | 0.7 | 0.65 | 0.65 | 55.16 | 0.65 | 1.11
| Analytical results for Gd in OREAS 353c (Indicative Value 1.54 ppm) | [15 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 1.51 | 1.6 | 1.1066 | 1.5 | 1.65 | 3 | [8 empty] | 1
| 1 | 2 | 1.46 | 1.6 | 1.1345 | 1.6 | 1.59 | 2 | [8 empty] | 6
| 1 | 3 | 1.5 | 1.6 | 1.0999 | 1.4 | 1.59 | 3 | [8 empty] | 16
| 1 | 4 | 1.39 | 1.6 | 1.1178 | 1.4 | 1.58 | < 2 | [8 empty] | 1.5367
| 1 | 5 | 1.41 | 1.6 | 1.1271 | 1.5 | 1.68 | < 2 | [8 empty] | 26
| 1 | 6 | 1.38 | 1.6 | 1.1463 | 1.5 | 1.64 | 3
| Mean |  | 1.4417 | 1.6 | 1.122 | 1.4833 | 1.6217 | 2.75
| Median |  | 1.435 | 1.6 | 1.1224 | 1.5 | 1.615 | 3
| Std Dev. |  | 0.056362 | 2.4324e-16 | 0.017412 | 0.075277 | 0.040702 | 0.5
| Rel.Std.Dev. |  | 3.909% | 1.52e-14% | 1.552% | 5.075% | 2.51% | 18.18%
| PDM3 |  | -6.182% | 4.121% | -26.98% | -3.471% | 5.531% | 78.96%
| Z-Score (Absolute) |  | 0.75 | 0.44 | 3.14 | 0.44 | 0.6 | 4.68
| Analytical results for Ge in OREAS 353c (Indicative Value < 1 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.12 | 1.13 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 0.5 | < 1 | < 1 | < 1 | [10 empty] | 1
| 1 | 2 | < 0.5 | < 1 | < 1 | < 1 | [10 empty] | 7
| 1 | 3 | < 0.5 | < 1 | < 1 | < 1 | [10 empty] | 16
| 1 | 4 | < 0.5 | < 1 | < 1 | < 1 | [10 empty] | < 1
| 1 | 5 | < 0.5 | < 1 | < 1 | < 1 | [10 empty] | 27
| 1 | 6 | < 0.5 | < 1 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Hf in OREAS 353c (Indicative Value 0.1 ppm) | [15 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.11 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 2 | [13 empty] | 1
| 1 | 2 | < 2 | 0.1 | [12 empty] | 8
| 1 | 3 | < 2 | 0.1 | [12 empty] | 16
| 1 | 4 | < 2 | 0.1 | [12 empty] | 0.1
| 1 | 5 | < 2 | [13 empty] | 28
| 1 | 6 | < 2 | 0.1
| Mean |  |  | 0.1
| Median |  |  | 0.1
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 353c (Indicative Value 0.23 ppm) | [15 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 0.22 | 0.3 | 0.2345 | 0.2 | 0.25 | < 2 | [8 empty] | 1
| 1 | 2 | 0.22 | 0.3 | 0.195 | 0.2 | 0.26 | < 2 | [8 empty] | 10
| 1 | 3 | 0.23 | 0.3 | 0.2065 | 0.2 | 0.26 | < 2 | [8 empty] | 16
| 1 | 4 | 0.23 | 0.3 | 0.2134 | 0.2 | 0.25 | < 2 | [8 empty] | 0.22648
| 1 | 5 | 0.22 | 0.3 | 0.2062 | 0.3 | 0.27 | < 2 | [8 empty] | 29
| 1 | 6 | 0.22 | 0.3 | 0.2085 | 0.2 | 0.27 | < 2
| Mean |  | 0.22333 | 0.3 | 0.21068 | 0.21667 | 0.26
| Median |  | 0.22 | 0.3 | 0.2075 | 0.2 | 0.26
| Std Dev. |  | 0.005164 | 0 | 0.013137 | 0.040825 | 0.0089443
| Rel.Std.Dev. |  | 2.312% | 0% | 6.236% | 18.84% | 3.44%
| PDM3 |  | -1.389% | 32.46% | -6.975% | -4.333% | 14.8%
| Z-Score (Absolute) |  | 0.44 | 1.41 | 0.75 | 0.6 | 0.44 | 18.27
| Analytical results for In in OREAS 353c (Certified Value 1.64 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.10 | 1.11 | 1.12 | 1.13 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 1.77 | 1.7 | 1.6 | 1.5 | 1.6529 | 1.5 | [8 empty] | 1
| 1 | 2 | 1.74 | 1.8 | 1.6 | 1.5 | 1.6457 | 1.6 | [8 empty] | #N/A
| 1 | 3 | 1.73 | 1.7 | 1.6 | 1.7 | 1.5683 | 1.6 | [8 empty] | 16
| 1 | 4 | 1.72 | 1.7 | 1.7 | 1.6 | 1.674 | 1.6 | [8 empty] | 1.641
| 1 | 5 | 1.85 | 1.7 | 1.6 | 1.5 | 1.692 | 1.5 | [8 empty] | 58
| 1 | 6 | 1.85 | 1.7 | 1.6 | 1.5 | 1.5843 | 1.5
| Mean |  | 1.7767 | 1.7167 | 1.6167 | 1.55 | 1.6362 | 1.55
| Median |  | 1.755 | 1.7 | 1.6 | 1.5 | 1.6493 | 1.55
| Std Dev. |  | 0.059217 | 0.040825 | 0.040825 | 0.083666 | 0.049447 | 0.054772
| Rel.Std.Dev. |  | 3.333% | 2.378% | 2.525% | 5.398% | 3.022% | 3.534%
| PDM3 |  | 8.265% | 4.609% | -1.485% | -5.547% | -0.2951% | -5.547%
| Z-Score (Absolute) |  | 1.33 | 0.8 | 0.09 | 0.67 | 0.09 | 0.67
| Analytical results for K in OREAS 353c (Indicative Value 0.042 wt.%) | [15 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 3
3
| 1 | 1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.058 | 0.058 | 0.14809 | < 0.1 | 0.06 | 0.15 | < 0.05 | < 0.05 | 0.0477 | 1
| 1 | 2 | < 0.1 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.058 | 0.05 | 0.14313 | < 0.1 |  | 0.15 | < 0.05 | < 0.05 | 0.0412 | 11
| 1 | 3 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | 0.058 | 0.1497 | < 0.1 |  | 0.14 | < 0.05 | < 0.05 | 0.0395 | 16
| 1 | 4 | < 0.1 | < 0.05 | < 0.05 | 0.05 | < 0.05 | 0.083 | 0.091 | 0.14335 | < 0.1 |  | 0.12 | < 0.05 | < 0.05 | 0.0323 | 0.04201
| 1 | 5 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | 0.066 | < 0.05 | 0.14757 | < 0.1 |  | 0.18 | < 0.05 | < 0.05 | 0.0376 | 30
| 1 | 6 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.05 | 0.05 | 0.14585 | < 0.1 |  | 0.16 | < 0.05 | < 0.05 | 0.0356
| Mean |  |  |  | 0.05 | 0.05 | 0.05 | 0.063 | 0.0614 | 0.14628 |  | 0.06 | 0.15 |  |  | 0.038983
| Median |  |  |  | 0.05 | 0.05 | 0.05 | 0.058 | 0.058 | 0.14671 |  | 0.06 | 0.15 |  |  | 0.03855
| Std Dev. | [5 empty] | 0 | 0.01253 | 0.017024 | 0.0026585 |  |  | 0.02 |  |  | 0.0052761
| Rel.Std.Dev. | [5 empty] | 0% | 19.89% | 27.73% | 1.817% |  |  | 13.33% |  |  | 13.53%
| PDM3 |  |  |  | 19.02% | 19.02% | 19.02% | 49.96% | 46.16% | 248.2% |  | 42.82% | 257.1% |  |  | -7.204%
| Z-Score (Absolute) |  | 0.24 | 0.85 | 0.67 | 0.67 | 0.49 | 0.53 | 0.47 | 4.45 | 0.24 | 0.68 | 4.61 | 0.85 | 0.85 | 0.24
| Analytical results for La in OREAS 353c (Indicative Value 9.96 ppm) | [15 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 9.33 | 10.5 | < 10 | 10.3 | 10.2 | 10 | [8 empty] | 1
| 1 | 2 | 9.06 | 10.3 | < 10 | 9.8 | 9.9 | 10 | [8 empty] | 13
| 1 | 3 | 9.12 | 10.4 | < 10 | 9.8 | 9.9 | 10 | [8 empty] | 16
| 1 | 4 | 9.35 | 10.7 | < 10 | 9.8 | 9.8 | 10 | [8 empty] | 9.9611
| 1 | 5 | 9.34 | 10.8 | < 10 | 10.2 | 9.7 | 10 | [8 empty] | 19
| 1 | 6 | 9.13 | 10.7 | < 10 | 9.9 | 10 | 10
| Mean |  | 9.2217 | 10.5667 |  | 9.9667 | 9.9167 | 10
| Median |  | 9.23 | 10.6 |  | 9.85 | 9.9 | 10
| Std Dev. |  | 0.13197 | 0.19664 |  | 0.22509 | 0.17224 | 0
| Rel.Std.Dev. |  | 1.431% | 1.861% |  | 2.258% | 1.737% | 0%
| PDM3 |  | -7.423% | 6.079% |  | 0.05577% | -0.4462% | 0.3904%
| Z-Score (Absolute) |  | 1.42 | 1.23 | 9.75 | 0.05 | 0.05 | 0.12
| Analytical results for Li in OREAS 353c (Indicative Value 5.04 ppm) | [15 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | < 5 | 2 | 10 | 1.4607 | 28 | 4 | 26 | 2.1749 | < 10 | 20 |  |  |  |  | 1
| 1 | 2 | < 5 | 2 | 10 | 1.4809 | < 10 | 6 |  | 2.144 | < 10 | 4 |  |  |  |  | 14
| 1 | 3 | < 5 | 2 | 10 | 1.3898 | < 10 | 5 |  | 2.1554 | < 10 | 7 |  |  |  |  | 16
| 1 | 4 | < 5 | 2 | < 10 | 1.4738 | < 10 | 18 |  | 2.1235 | < 10 | 3 |  |  |  |  | 5.0372
| 1 | 5 | < 5 | 2 | 10 | 1.4039 | < 10 | 10 |  | 2.1489 | < 10 | 4 |  |  |  |  | 20
| 1 | 6 | < 5 | 2 | 10 | 1.3953 | < 10 | 4 |  | 2.1881 | < 10 | 15
| Mean |  |  | 2 | 10 | 1.4341 | 28 | 7.8333 | 26 | 2.1558 |  | 8.8333
| Median |  |  | 2 | 10 | 1.4323 | 28 | 5.5 | 26 | 2.1521 |  | 5.5
| Std Dev. |  |  | 0 | 0 | 0.04209 |  | 5.4559 |  | 0.02297 |  | 7.0261
| Rel.Std.Dev. |  |  | 0% | 0% | 2.935% |  | 69.65% |  | 1.065% |  | 79.54%
| PDM3 |  |  | -60.3% | 98.52% | -71.53% | 455.9% | 55.51% | 416.2% | -57.2% |  | 75.36%
| Z-Score (Absolute) |  | 0.79 | 0.89 | 0.56 | 1.01 | 0.49 | 0.29 | 3.96 | 0.86 | 0.29 | 0.49
| Analytical results for Lu in OREAS 353c (Indicative Value 0.064 ppm) | [15 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 3
3
| 1 | 1 | 0.05 | 0.08 | 0.0519 | 0.07 | < 2 | [9 empty] | 1
| 1 | 2 | 0.05 | 0.1 | 0.0557 | 0.06 | < 2 | [9 empty] | 3
| 1 | 3 | 0.05 | 0.07 | 0.0482 | 0.07 | < 2 | [9 empty] | 16
| 1 | 4 | 0.05 | 0.09 | 0.0504 | 0.07 | < 2 | [9 empty] | 0.063642
| 1 | 5 | 0.06 | 0.08 | 0.0534 | 0.07 | < 2 | [9 empty] | 21
| 1 | 6 | 0.05 | 0.09 | 0.0478 | 0.06 | < 2
| Mean |  | 0.051667 | 0.085 | 0.051233 | 0.066667
| Median |  | 0.05 | 0.085 | 0.05115 | 0.07
| Std Dev. |  | 0.0040825 | 0.010488 | 0.0030585 | 0.005164
| Rel.Std.Dev. |  | 7.902% | 12.34% | 5.97% | 7.746%
| PDM3 |  | -18.82% | 33.56% | -19.5% | 4.753%
| Z-Score (Absolute) |  | 0.66 | 0.8 | 0.67 | 0 | 40.78
| Analytical results for Mg in OREAS 353c (Certified Value 0.156 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 3
3
| 1 | 1 | 0.15 | 0.16 | 0.151 | 0.121 | 0.157 | 0.157 | 0.169 | 0.1263 | 0.152 | 0.15 | 0.16 | 0.15012 | 0.16 | 0.1618 | 1
| 1 | 2 | 0.14 | 0.16 | 0.157 | 0.121 | 0.151 | 0.157 | 0.163 | 0.12553 | 0.166 | 0.15 | 0.15 | 0.14836 | 0.15 | 0.1674 | #N/A
| 1 | 3 | 0.15 | 0.16 | 0.151 | 0.121 | 0.163 | 0.157 | 0.163 | 0.12861 | 0.166 | 0.15 | 0.15 | 0.14857 | 0.16 | 0.1633 | 16
| 1 | 4 | 0.15 | 0.16 | 0.151 | 0.139 | 0.169 | 0.157 | 0.163 | 0.12629 | 0.164 | 0.15 | 0.15 | 0.15347 | 0.15 | 0.1626 | 0.15585
| 1 | 5 | 0.16 | 0.15 | 0.157 | 0.109 | 0.163 | 0.169 | 0.163 | 0.12801 | 0.159 | 0.15 | 0.15 | 0.15125 | 0.15 | 0.1645 | 59
| 1 | 6 | 0.14 | 0.16 | 0.157 | 0.121 | 0.157 | 0.157 | 0.151 | 0.12643 | 0.158 | 0.15 | 0.15 | 0.15214 | 0.15 | 0.1575
| Mean |  | 0.14833 | 0.15833 | 0.154 | 0.122 | 0.16 | 0.159 | 0.162 | 0.12686 | 0.16083 | 0.15 | 0.15167 | 0.15065 | 0.15333 | 0.16285
| Median |  | 0.15 | 0.16 | 0.154 | 0.121 | 0.16 | 0.157 | 0.163 | 0.12637 | 0.1615 | 0.15 | 0.15 | 0.15069 | 0.15 | 0.16295
| Std Dev. |  | 0.0075277 | 0.0040825 | 0.0032863 | 0.0096125 | 0.0062929 | 0.004899 | 0.0058992 | 0.0011806 | 0.0055287 | 0 | 0.0040825 | 0.002019 | 0.005164 | 0.0032685
| Rel.Std.Dev. |  | 5.075% | 2.578% | 2.134% | 7.879% | 3.933% | 3.081% | 3.641% | 0.9306% | 3.438% | 0% | 2.692% | 1.34% | 3.368% | 2.007%
| PDM3 |  | -4.824% | 1.592% | -1.188% | -21.72% | 2.661% | 2.02% | 3.945% | -18.6% | 3.196% | -3.755% | -2.685% | -3.336% | -1.616% | 4.49%
| Z-Score (Absolute) |  | 0.67 | 0.61 | 0.03 | 4.11 | 0.81 | 0.68 | 1.06 | 3.44 | 0.94 | 0.46 | 0.24 | 0.37 | 0.03 | 1.2
| Analytical results for Mn in OREAS 353c (Certified Value 0.137 wt.%) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 3
3
| 1 | 1 | 0.1352 | 0.136 | 0.132 | 0.139 | 0.132 | 0.132 | 0.139 | 0.13578 | 0.1431 | 0.1435 | 0.2 | 0.13915 | 0.141 | 0.14 | 1
| 1 | 2 | 0.1314 | 0.137 | 0.139 | 0.139 | 0.132 | 0.132 | 0.139 | 0.13572 | 0.1466 | 0.1422 |  | 0.13518 | 0.138 | 0.1424 | #N/A
| 1 | 3 | 0.1346 | 0.134 | 0.132 | 0.139 | 0.139 | 0.132 | 0.139 | 0.13582 | 0.1469 | 0.1442 |  | 0.13906 | 0.143 | 0.1425 | 16
| 1 | 4 | 0.1322 | 0.136 | 0.132 | 0.132 | 0.139 | 0.132 | 0.132 | 0.13486 | 0.146 | 0.1388 | 0.3 | 0.1321 | 0.139 | 0.1381 | 0.13747
| 1 | 5 | 0.1329 | 0.133 | 0.139 | 0.139 | 0.139 | 0.132 | 0.139 | 0.13766 | 0.1425 | 0.1427 |  | 0.13512 | 0.139 | 0.1437 | 60
| 1 | 6 | 0.1334 | 0.136 | 0.132 | 0.139 | 0.139 | 0.132 | 0.139 | 0.13513 | 0.1409 | 0.1444 | 0.2 | 0.13343 | 0.139 | 0.1363
| Mean |  | 0.13328 | 0.13533 | 0.13433 | 0.13783 | 0.13667 | 0.132 | 0.13783 | 0.13583 | 0.14433 | 0.14263 | 0.23333 | 0.13567 | 0.13983 | 0.1405
| Median |  | 0.13315 | 0.136 | 0.132 | 0.139 | 0.139 | 0.132 | 0.139 | 0.13575 | 0.14455 | 0.1431 | 0.2 | 0.13515 | 0.139 | 0.1412
| Std Dev. |  | 0.0014345 | 0.0015055 | 0.0036148 | 0.0028577 | 0.0036148 | 0 | 0.0028577 | 0.00098039 | 0.0024969 | 0.0020598 | 0.057735 | 0.002895 | 0.0018348 | 0.002881
| Rel.Std.Dev. |  | 1.076% | 1.112% | 2.691% | 2.073% | 2.645% | 0% | 2.073% | 0.7218% | 1.73% | 1.444% | 24.74% | 2.134% | 1.312% | 2.051%
| PDM3 |  | -3.046% | -1.555% | -2.282% | 0.2635% | -0.5851% | -3.98% | 0.2635% | -1.194% | 4.992% | 3.755% | 69.73% | -1.309% | 1.718% | 2.203%
| Z-Score (Absolute) |  | 0.92 | 0.39 | 0.67 | 0.34 | 0 | 1.35 | 0.34 | 0.26 | 1.96 | 1.52 | NA | 0.3 | 0.79 | 0.96
Results from laboratory 11 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Mo in OREAS 353c (Certified Value 11 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | [6 empty] | 1
1
| 1 | 1 | 8.87 | 12 | 12 | 10 |  | 12.1453 | 10 | 13 | [6 empty] | 1
| 1 | 2 | 8.59 | 11 | 12 | 11 |  | 12.0178 | 10 | 11 | [6 empty] | #N/A
| 1 | 3 | 8.35 | 12 | < 10 | 11 |  | 11.7403 | 10 | 12 | [6 empty] | 16
| 1 | 4 | 9.41 | 12 | 12 | 12 | 60 | 11.9947 | 10 | 9 | [6 empty] | 11.01
| 1 | 5 | 8.76 | 12 | 14 | 11 |  | 11.7605 | 10 | 10 | [6 empty] | 61
| 1 | 6 | 8.87 | 12 | < 10 | 11 |  | 11.9107 | 10 | 11
| Mean |  | 8.8083 | 11.8333 | 12.5 | 11 | 60 | 11.9282 | 10 | 11
| Median |  | 8.815 | 12 | 12 | 11 | 60 | 11.9527 | 10 | 11
| Std Dev. |  | 0.35488 | 0.40825 | 1 | 0.63246 |  | 0.15705 | 0 | 1.4142
| Rel.Std.Dev. |  | 4.029% | 3.45% | 8% | 5.75% |  | 1.317% | 0% | 12.86%
| PDM3 |  | -20% | 7.478% | 13.53% | -0.09066% | 445% | 8.34% | -9.173% | -0.09066%
| Z-Score (Absolute) |  | 1.53 | 0.58 | 0.7 | 0 | 34.27 | 0.65 | 0.7 | 0
| Analytical results for Nb in OREAS 353c (Indicative Value < 5 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.11 | 1.12 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | < 5 | 14 | < 2 | 6 | < 5 | [9 empty] | 1
| 1 | 2 | < 5 | 10 | < 2 | < 5 | < 5 | [9 empty] | 5
| 1 | 3 | < 5 |  | < 2 | < 5 | < 5 | [9 empty] | 16
| 1 | 4 | < 5 | 17 | < 2 | 7 | < 5 | [9 empty] | < 5
| 1 | 5 | < 5 | 11 | < 2 | < 5 | < 5 | [9 empty] | 22
| 1 | 6 | < 5 | 13 | < 2 | 6 | < 5
| Mean |  |  | 13 |  | 6.3333
| Median |  |  | 13 |  | 6
| Std Dev. |  |  | 2.7386 |  | 0.57735
| Rel.Std.Dev. |  |  | 21.07% |  | 9.116%
PDM3
| Z-Score (Absolute) |  | 0 | 4.72 | 0.67 | 0.86 | 0
| Analytical results for Nd in OREAS 353c (Indicative Value 7.09 ppm) | [15 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 5.66 | 7.1 | 5.7123 | 6.9 | 7 | 9 | [8 empty] | 1
| 1 | 2 | 5.64 | 7.2 | 5.8262 | 7.2 | 8 | 6 | [8 empty] | 6
| 1 | 3 | 5.82 | 7.4 | 5.6905 | 7 | 8 | 8 | [8 empty] | 16
| 1 | 4 | 5.73 | 7.1 | 5.5637 | 6.9 | 7 | 5 | [8 empty] | 7.0917
| 1 | 5 | 5.58 | 7 | 5.5988 | 6.8 | 7 | 5 | [8 empty] | 23
| 1 | 6 | 5.85 | 7.2 | 5.549 | 7.4 | 7 | 8
| Mean |  | 5.7133 | 7.1667 | 5.6568 | 7.0333 | 7.3333 | 6.8333
| Median |  | 5.695 | 7.15 | 5.6446 | 6.95 | 7 | 7
| Std Dev. |  | 0.10614 | 0.13663 | 0.10636 | 0.22509 | 0.5164 | 1.7224
| Rel.Std.Dev. |  | 1.858% | 1.906% | 1.88% | 3.2% | 7.042% | 25.21%
| PDM3 |  | -19.44% | 1.058% | -20.23% | -0.8226% | 3.408% | -3.643%
| Z-Score (Absolute) |  | 2.6 | 0.5 | 2.72 | 0.21 | 0.85 | 0.21
| Analytical results for Ni in OREAS 353c (Indicative Value 15.7 ppm) | [15 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 1
1
| 1 | 1 | 23 | < 20 | < 20 | < 20 | 40 | < 20 | 20 | 21 | 21 | < 20 | < 10 |  |  |  | 1
| 1 | 2 | 23 | < 20 | < 20 | < 20 | 40 | < 20 | < 20 | 31 | 30 | < 20 | 10 |  |  |  | 18
| 1 | 3 | 21 | < 20 | < 20 | < 20 | 20 | < 20 | 20 | 20 | 20 | < 20 | 14 |  |  |  | 16
| 1 | 4 | 21 | 20 | < 20 | 40 | 40 | < 20 | < 20 | 14 | 28 | < 20 | 12 |  |  |  | 15.7083
| 1 | 5 | 21 | 20 | < 20 | < 20 | 40 | 40 | < 20 | 18 | 23 | < 20 | 18 |  |  |  | 24
| 1 | 6 | 23 | < 20 | < 20 | < 20 | 40 | 70 | < 20 | 23 | 23 | < 20 | 11
| Mean |  | 22 | 20 |  | 40 | 36.6667 | 55 | 20 | 21.1667 | 24.1667 |  | 13
| Median |  | 22 | 20 |  | 40 | 40 | 55 | 20 | 20.5 | 23 |  | 12
| Std Dev. |  | 1.0954 | 0 |  |  | 8.165 | 21.2132 | 0 | 5.7067 | 3.9707 |  | 3.1623
| Rel.Std.Dev. |  | 4.979% | 0% |  |  | 22.27% | 38.57% | 0% | 26.96% | 16.43% |  | 24.33%
| PDM3 |  | 40.05% | 27.32% |  | 154.6% | 133.4% | 250.1% | 27.32% | 34.75% | 53.85% |  | -17.24%
| Z-Score (Absolute) |  | 0.94 | 0.22 | 0.67 | 0 | 2.92 | 1.35 | 0.22 | 0.83 | 1.24 | 0.67 | 0.45
| Analytical results for P in OREAS 353c (Indicative Value 0.021 wt.%) | [15 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.13 | 1.14 | [9 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | [9 empty] | 3
3
| 1 | 1 | 0.02 | 0.02 | 0.017 | 0.03 | 0.018 | [9 empty] | 1
| 1 | 2 | 0.02 | 0.02 | 0.014 | 0.03 | 0.0177 | [9 empty] | 19
| 1 | 3 | 0.02 | 0.02 | 0.016 | 0.03 | 0.0179 | [9 empty] | 16
| 1 | 4 | 0.02 | 0.02 | 0.014 | 0.03 | 0.0167 | [9 empty] | 0.020587
| 1 | 5 | 0.02 | 0.02 | 0.016 | 0.03 | 0.0165 | [9 empty] | 25
| 1 | 6 | 0.02 | 0.02 | 0.018 | 0.03 | 0.0158
| Mean |  | 0.02 | 0.02 | 0.015833 | 0.03 | 0.0171
| Median |  | 0.02 | 0.02 | 0.016 | 0.03 | 0.0172
| Std Dev. |  | 0 | 0 | 0.0016021 | 0 | 0.00089666
| Rel.Std.Dev. |  | 0% | 0% | 10.12% | 0% | 5.244%
| PDM3 |  | -2.85% | -2.85% | -23.09% | 45.73% | -16.94%
| Z-Score (Absolute) |  | 0 | 0 | 0.97 | 2.33 | 0.67
| Analytical results for Pb in OREAS 353c (Indicative Value 71.33 wt.%) | [15 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS*LW | PF*OES/MS |  |  | 2
[13 empty] | LWg |  |  |  | 2
| 1 | 1 | 72.45 | 45.4 | > 30 | > 30 | > 30 | > 30 | > 30 | 70.6584 | > 10 | > 70 | 73.5068 | 66.1 |  |  | 1
| 1 | 2 | 72.01 | 45.7 | > 30 | > 30 | > 30 | > 30 | > 30 | 70.7128 | > 10 | > 70 | 73.5065 | 65.2 |  |  | 20
| 1 | 3 | 72.35 | 45.4 | > 30 | > 30 | > 30 | > 30 | > 30 | 70.2433 | > 10 | > 70 | 74.1262 | 70.3 |  |  | 16
| 1 | 4 | 71.99 | 44.2 | > 30 | > 30 | > 30 | > 30 | > 30 | 70.5416 | > 10 | > 70 | 73.8372 | 71 |  |  | 71.3327
| 1 | 5 | 72.15 | 43.8 | > 30 | > 30 | > 30 | > 30 | > 30 | 70.3834 | > 10 | > 70 | 74.3935 | 67 |  |  | 26
| 1 | 6 | 72.51 | 44.9 | > 30 | > 30 | > 30 | > 30 | > 30 | 70.5774 | > 10 | > 70 | 74.3386 | 72.1
| Mean |  | 72.2433 | 44.9 | [5 empty] | 70.5195 |  |  | 73.9515 | 68.6167
| Median |  | 72.25 | 45.15 | [5 empty] | 70.5595 |  |  | 73.9817 | 68.65
| Std Dev. |  | 0.22474 | 0.75366 | [5 empty] | 0.17624 |  |  | 0.39613 | 2.8729
| Rel.Std.Dev. |  | 0.3111% | 1.679% | [5 empty] | 0.2499% |  |  | 0.5357% | 4.187%
| PDM3 |  | 1.277% | -37.06% | [5 empty] | -1.14% |  |  | 3.671% | -3.808%
| Z-Score (Absolute) |  | 0.61 | 9.08 | NA | NA | NA | NA | NA | 0 | NA | NA | 1.22 | 0.67
| Analytical results for Pr in OREAS 353c (Indicative Value 1.89 ppm) | [15 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.11 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 2.01 | 1.94 | 1.8 | 1.87 | 4 | [9 empty] | 1
| 1 | 2 | 2.1 | 1.96 | 1.9 | 1.81 | 3 | [9 empty] | 7
| 1 | 3 | 1.97 | 1.86 | 1.8 | 1.9 | 4 | [9 empty] | 16
| 1 | 4 | 2.07 | 1.96 | 1.9 | 1.8 | 3 | [9 empty] | 1.8915
| 1 | 5 | 1.97 | 1.86 | 1.8 | 1.8 | 3 | [9 empty] | 27
| 1 | 6 | 1.91 | 1.89 | 1.8 | 1.8 | 4
| Mean |  | 2.005 | 1.9117 | 1.8333 | 1.83 | 3.5
| Median |  | 1.99 | 1.915 | 1.8 | 1.805 | 3.5
| Std Dev. |  | 0.070356 | 0.047504 | 0.05164 | 0.043818 | 0.54772
| Rel.Std.Dev. |  | 3.509% | 2.485% | 2.817% | 2.394% | 15.65%
| PDM3 |  | 6.001% | 1.066% | -3.075% | -3.251% | 85.04%
| Z-Score (Absolute) |  | 0.77 | 0 | 0.65 | 0.67 | 13.11
| Analytical results for Rb in OREAS 353c (Indicative Value 1.35 ppm) | [15 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [7 empty] | 2
2
| 1 | 1 | 1.22 | 1.4 | 0.81566 | 2.1 | 1.4 | 1.3149 | < 2 | [7 empty] | 1
| 1 | 2 | 1.31 | 1.6 | 0.7583 | 1.6 | 1.2 | 1.2616 | < 2 | [7 empty] | 8
| 1 | 3 | 1.21 | 1.4 | 0.85992 | 1.7 | 1.4 | 1.4563 | < 2 | [7 empty] | 16
| 1 | 4 | 1.32 | 1.3 | 0.73848 | 1.8 | 1.4 | 1.2766 | < 2 | [7 empty] | 1.3509
| 1 | 5 | 1.31 | 1.3 | 0.79582 | 1.9 | 1.7 | 1.2731 | < 2 | [7 empty] | 28
| 1 | 6 | 1.29 | 1.3 | 0.81255 | 1.6 | 1.3 | 1.478 | < 2
| Mean |  | 1.2767 | 1.3833 | 0.79679 | 1.7833 | 1.4 | 1.3434
| Median |  | 1.3 | 1.35 | 0.80418 | 1.75 | 1.4 | 1.2958
| Std Dev. |  | 0.048854 | 0.1169 | 0.043525 | 0.19408 | 0.16733 | 0.097724
| Rel.Std.Dev. |  | 3.827% | 8.451% | 5.463% | 10.88% | 11.95% | 7.274%
| PDM3 |  | -5.492% | 2.404% | -41.02% | 32.02% | 3.638% | -0.551%
| Z-Score (Absolute) |  | 0.67 | 0.4 | 5.52 | 4.44 | 0.57 | 0 | 3.47
| Analytical results for Re in OREAS 353c (Indicative Value < 0.1 ppm) | [15 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.12 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | [11 empty] | 2
2
| 1 | 1 | < 1 | < 0.1 | < 0.1 | [11 empty] | 1
| 1 | 2 | < 1 | < 0.1 | < 0.1 | [11 empty] | 9
| 1 | 3 | < 1 | < 0.1 | < 0.1 | [11 empty] | 16
| 1 | 4 | < 1 | < 0.1 | < 0.1 | [11 empty] | < 0.1
| 1 | 5 | < 1 | < 0.1 | < 0.1 | [11 empty] | 29
| 1 | 6 | < 1 | < 0.1 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 353c (Certified Value 14.69 wt.%) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS |  |  | 2
3
| 1 | 1 | 14.2 | 14.75 | 14.8 | 15.55 | 15 | 14.6 | 15.6536 | 14.62 | 14.49 | 14.776 | 14.4 | 13.8924 |  |  | 1
| 1 | 2 | 14.1 | 14.75 | 14.7 | 14.95 | 15.15 | 14.7 | 15.6393 | 14.63 | 13.83 | 14.5725 | 14.4 | 14.0724 |  |  | #N/A
| 1 | 3 | 14.2 | 14.55 | 14.65 | 15.8 | 14.95 | 15.05 | 15.6329 | 14.37 | 14.22 | 14.9123 | 14.8 | 14.0036 |  |  | 16
| 1 | 4 | 14.1 | 14.75 | 14.65 | 16 | 14.85 | 14.7 | 15.5701 | 14.06 | 13.88 | 14.9533 | 14.4 | 13.6318 |  |  | 14.6942
| 1 | 5 | 14.5 | 14.9 | 14.7 | 15.7 | 14.85 | 14.6 | 15.6389 | 14.6 | 14.16 | 14.92 | 14.7 | 14.1583 |  |  | 62
| 1 | 6 | 14.3 | 14.85 | 14.7 | 15.6 | 14.95 | 14.75 | 15.5826 | 14.4 | 14.19 | 14.545 | 14.5 | 13.2808
| Mean |  | 14.2333 | 14.7583 | 14.7 | 15.6 | 14.9583 | 14.7333 | 15.6196 | 14.4467 | 14.1283 | 14.7798 | 14.5333 | 13.8399
| Median |  | 14.2 | 14.75 | 14.7 | 15.65 | 14.95 | 14.7 | 15.6359 | 14.5 | 14.175 | 14.8441 | 14.45 | 13.948
| Std Dev. |  | 0.15055 | 0.12007 | 0.054772 | 0.35637 | 0.11143 | 0.16633 | 0.034405 | 0.22124 | 0.24277 | 0.1819 | 0.17512 | 0.3289
| Rel.Std.Dev. |  | 1.058% | 0.8136% | 0.3726% | 2.284% | 0.7449% | 1.129% | 0.2203% | 1.531% | 1.718% | 1.231% | 1.205% | 2.376%
| PDM3 |  | -3.137% | 0.4361% | 0.03916% | 6.164% | 1.797% | 0.266% | 6.297% | -1.685% | -3.851% | 0.5825% | -1.095% | -5.814%
| Z-Score (Absolute) |  | 1.27 | 0.11 | 0.04 | 2.33 | 0.64 | 0.04 | 2.38 | 0.71 | 1.55 | 0.17 | 0.48 | 2.31
| Analytical results for Sb in OREAS 353c (Certified Value 0.28 wt.%) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | [6 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.2442 | 0.284 | 0.30406 | 0.3077 | 0.28269 | 0.31324 | 0.27828 | 0.298 | [6 empty] | 1
| 1 | 2 | 0.2351 | 0.282 | 0.30413 | 0.3139 | 0.27119 | 0.25768 | 0.26823 | 0.293 | [6 empty] | #N/A
| 1 | 3 | 0.2378 | 0.288 | 0.30304 | 0.3031 | 0.28218 | 0.25312 | 0.27529 | 0.296 | [6 empty] | 16
| 1 | 4 | 0.2433 | 0.282 | 0.30392 | 0.293 | 0.27134 | 0.25101 | 0.26597 | 0.285 | [6 empty] | 0.27979
| 1 | 5 | 0.2383 | 0.288 | 0.30557 | 0.2914 | 0.27644 | 0.25833 | 0.27645 | 0.299 | [6 empty] | 63
| 1 | 6 | 0.2412 | 0.286 | 0.30698 | 0.2852 | 0.2804 | 0.25953 | 0.27245 | 0.293
| Mean |  | 0.23998 | 0.285 | 0.30462 | 0.29905 | 0.27737 | 0.26549 | 0.27278 | 0.294
| Median |  | 0.23975 | 0.285 | 0.3041 | 0.29805 | 0.27842 | 0.25801 | 0.27387 | 0.2945
| Std Dev. |  | 0.0035131 | 0.0027568 | 0.0014151 | 0.010943 | 0.0052161 | 0.023624 | 0.0048419 | 0.0050596
| Rel.Std.Dev. |  | 1.464% | 0.9673% | 0.4645% | 3.659% | 1.881% | 8.899% | 1.775% | 1.721%
| PDM3 |  | -14.23% | 1.864% | 8.875% | 6.885% | -0.8623% | -5.111% | -2.504% | 5.08%
| Z-Score (Absolute) |  | 1.95 | 0.18 | 1.11 | 0.84 | 0.18 | 0.74 | 0.4 | 0.61
| Analytical results for Sc in OREAS 353c (Indicative Value < 5 ppm) | [15 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.12 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | < 5 | < 20 | < 5 | < 20 | < 10 | < 2 | [8 empty] | 1
| 1 | 2 | < 5 | < 20 | < 5 | < 20 | < 10 | < 2 | [8 empty] | 24
| 1 | 3 | < 5 | < 20 | < 5 | < 20 | < 10 | < 2 | [8 empty] | 16
| 1 | 4 | < 5 | < 20 | < 5 | < 20 | < 10 | < 2 | [8 empty] | < 5
| 1 | 5 | < 5 | < 20 | < 5 | < 20 | < 10 | < 2 | [8 empty] | 30
| 1 | 6 | < 5 | < 20 | < 5 | < 20 | < 10 | < 2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Se in OREAS 353c (Indicative Value < 20 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.11 | 1.12 | 1.13 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [10 empty] | 1
1
| 1 | 1 | < 20 | 21 | < 20 | 5 | [10 empty] | 1
| 1 | 2 | < 20 | 20 | < 20 | < 5 | [10 empty] | 10
| 1 | 3 | < 20 |  | < 20 | 6 | [10 empty] | 16
| 1 | 4 | < 20 | 28 | < 20 | 6 | [10 empty] | < 20
| 1 | 5 | < 20 |  | < 20 | < 5 | [10 empty] | 19
| 1 | 6 | < 20 |  | < 20 | < 5
| Mean |  |  | 23 |  | 5.6667
| Median |  |  | 21 |  | 6
| Std Dev. |  |  | 4.3589 |  | 0.57735
| Rel.Std.Dev. |  |  | 18.95% |  | 10.19%
PDM3
| Z-Score (Absolute) |  | 0 | 2.96 | 0 | 1.35
| Analytical results for Si in OREAS 353c (Certified Value 1.48 wt.%) | [15 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 |  |  | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 2
3
| 1 | 1 | 1.45 | 1.7 | 1.5 | 1.5 | 1.68 | 1.5 | 1.59 | 1.5962 | 1.435 | 1.5 | 1.5258 | 1.5 |  |  | 1
| 1 | 2 | 1.42 | 1.7 | 1.5 | 1.4 | 1.5 | 1.5 | 1.5 | 1.5886 | 1.511 | 1.5 | 1.4628 | 1.5 |  |  | #N/A
| 1 | 3 | 1.45 | 1.7 | 1.5 | 1.4 | 1.59 | 1.5 | 1.5 | 1.5822 | 1.511 | 1.5 | 1.5185 | 1.5 |  |  | 16
| 1 | 4 | 1.43 | 1.7 | 1.5 | 1.5 | 1.59 | 1.5 | 1.4 | 1.5823 | 1.493 | 1.5 | 1.5645 | 1.5 |  |  | 1.4824
| 1 | 5 | 1.42 | 1.7 | 1.5 | 1.5 | 1.68 | 1.5 | 1.5 | 1.5951 | 1.433 | 1.5 | 1.5306 | 1.5 |  |  | 64
| 1 | 6 | 1.44 | 1.7 | 1.5 | 1.4 | 1.68 | 1.5 | 1.5 | 1.5921 | 1.387 | 1.5 | 1.4475 | 1.5
| Mean |  | 1.435 | 1.7 | 1.5 | 1.45 | 1.62 | 1.5 | 1.4983 | 1.5894 | 1.4617 | 1.5 | 1.5083 | 1.5
| Median |  | 1.435 | 1.7 | 1.5 | 1.45 | 1.635 | 1.5 | 1.5 | 1.5904 | 1.464 | 1.5 | 1.5222 | 1.5
| Std Dev. |  | 0.013784 | 2.4324e-16 | 0 | 0.054772 | 0.073485 | 0 | 0.060139 | 0.0061347 | 0.050906 | 0 | 0.044358 | 0
| Rel.Std.Dev. |  | 0.9606% | 1.431e-14% | 0% | 3.777% | 4.536% | 0% | 4.014% | 0.386% | 3.483% | 0% | 2.941% | 0%
| PDM3 |  | -3.199% | 14.68% | 1.186% | -2.187% | 9.281% | 1.186% | 1.074% | 7.219% | -1.4% | 1.186% | 1.746% | 1.186%
| Z-Score (Absolute) |  | 1.82 | 5.84 | 0.06 | 1.41 | 3.55 | 0.06 | 0.06 | 2.65 | 1.05 | 0.06 | 0.3 | 0.06
| Analytical results for Sm in OREAS 353c (Certified Value 1.42 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.08 | 1.11 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 1.22 | 1.3142 | 1.7 | 1.5 | 2 | [9 empty] | 1
| 1 | 2 | 1.32 | 1.2739 | 1.5 | 1.5 | < 2 | [9 empty] | 11
| 1 | 3 | 1.23 | 1.2236 | 1.5 | 1.4 | 3 | [9 empty] | 16
| 1 | 4 | 1.17 | 1.2035 | 1.4 | 1.6 | < 2 | [9 empty] | 1.4213
| 1 | 5 | 1.2 | 1.2981 | 1.6 | 1.5 | 2 | [9 empty] | 65
| 1 | 6 | 1.29 | 1.2962 | 1.4 | 1.4 | 2
| Mean |  | 1.2383 | 1.2682 | 1.5167 | 1.4833 | 2.25
| Median |  | 1.225 | 1.2851 | 1.5 | 1.5 | 2
| Std Dev. |  | 0.056362 | 0.044725 | 0.1169 | 0.075277 | 0.5
| Rel.Std.Dev. |  | 4.551% | 3.527% | 7.708% | 5.075% | 22.22%
| PDM3 |  | -12.87% | -10.77% | 6.709% | 4.363% | 58.3%
| Z-Score (Absolute) |  | 0.77 | 0.67 | 0.1 | 0 | 1.1
| Analytical results for Sn in OREAS 353c (Indicative Value 15.3 ppm) | [15 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 1
1
| 1 | 1 | < 10 | < 50 | 18 | < 100 | 17 | 21 | [8 empty] | 1
| 1 | 2 | < 10 | < 50 | 17 | < 100 | 17 | < 20 | [8 empty] | 14
| 1 | 3 | < 10 | < 50 | 17 | < 100 | 16 | < 20 | [8 empty] | 16
| 1 | 4 | < 10 | < 50 | 18 | < 100 | 16 | < 20 | [8 empty] | 15.2889
| 1 | 5 | < 10 | < 50 | 17 | < 100 | 16 | 23 | [8 empty] | 20
| 1 | 6 | < 10 | < 50 | 17 | < 100 | 16 | < 20
| Mean |  |  |  | 17.3333 |  | 16.3333 | 22
| Median |  |  |  | 17 |  | 16 | 22
| Std Dev. |  |  |  | 0.5164 |  | 0.5164 | 1.4142
| Rel.Std.Dev. |  |  |  | 2.979% |  | 3.162% | 6.428%
| PDM3 |  |  |  | 13.37% |  | 6.831% | 43.9%
| Z-Score (Absolute) |  | 1.45 | 1 | 0.06 | 4.07 | 0.06 | 0.35
| Analytical results for Sr in OREAS 353c (Certified Value < 10 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.12 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 1
1
| 1 | 1 | 7.63 | < 20 | < 10 | < 20 | < 10 | 26 | [8 empty] | 1
| 1 | 2 | 7.58 | < 20 | < 10 | < 20 | < 10 | < 5 | [8 empty] | 3
| 1 | 3 | 7.14 | < 20 | < 10 | < 20 | < 10 | 73 | [8 empty] | 16
| 1 | 4 | 7.36 | < 20 | < 10 | < 20 | < 10 | < 5 | [8 empty] | < 10
| 1 | 5 | 7.78 | < 20 | < 10 | < 20 | < 10 | 68 | [8 empty] | 66
| 1 | 6 | 7.27 | < 20 | < 10 | < 20 | < 10 | 16
| Mean |  | 7.46 |  |  |  |  | 45.75
| Median |  | 7.47 |  |  |  |  | 47
| Std Dev. |  | 0.24257 |  |  |  |  | 28.941
| Rel.Std.Dev. |  | 3.252% |  |  |  |  | 63.26%
PDM3
| Z-Score (Absolute) |  | 0.34 | 0.34 | 1.01 | 0.34 | 1.01 | 6.1
| Analytical results for Ta in OREAS 353c (Indicative Value 0.21 ppm) | [15 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 1.36 | 0.2 |  | 0.3 | < 0.1 | < 0.5 | < 10 | [7 empty] | 1
| 1 | 2 | 1.23 | 0.2 | 0.2 | 0.3 | < 0.1 | < 0.5 | < 10 | [7 empty] | 15
| 1 | 3 | 1.27 | 0.2 |  | 0.3 | < 0.1 | < 0.5 | < 10 | [7 empty] | 16
| 1 | 4 | 1.27 | 0.2 | 0.1 | 0.4 | < 0.1 | < 0.5 | < 10 | [7 empty] | 0.21111
| 1 | 5 | 1.33 | 0.2 | 0.2 | 0.2 | < 0.1 | < 0.5 | < 10 | [7 empty] | 21
| 1 | 6 | 1.4 | 0.2 | 0.1 | 0.2 | < 0.1 | < 0.5 | < 10
| Mean |  | 1.31 | 0.2 | 0.15 | 0.28333
| Median |  | 1.3 | 0.2 | 0.15 | 0.3
| Std Dev. |  | 0.064187 | 3.0405e-17 | 0.057735 | 0.075277
| Rel.Std.Dev. |  | 4.9% | 1.52e-14% | 38.49% | 26.57%
| PDM3 |  | 520.5% | -5.263% | -28.95% | 34.21%
| Z-Score (Absolute) |  | 7.15 | 0.34 | 0.67 | 0.22 | 1.35 | 0 | 32.03
| Analytical results for Tb in OREAS 353c (Indicative Value 0.19 ppm) | [15 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 0.16 | 0.25 | 0.115 | 0.2 | 0.21 | < 2 | [8 empty] | 1
| 1 | 2 | 0.15 | 0.25 | 0.1138 | 0.2 | 0.23 | < 2 | [8 empty] | 4
| 1 | 3 | 0.16 | 0.24 | 0.1172 | 0.2 | 0.23 | < 2 | [8 empty] | 16
| 1 | 4 | 0.17 | 0.25 | 0.1163 | 0.2 | 0.22 | < 2 | [8 empty] | 0.18827
| 1 | 5 | 0.17 | 0.24 | 0.1278 | 0.2 | 0.21 | < 2 | [8 empty] | 22
| 1 | 6 | 0.15 | 0.25 | 0.1469 | 0.2 | 0.2 | < 2
| Mean |  | 0.16 | 0.24667 | 0.12283 | 0.2 | 0.21667
| Median |  | 0.16 | 0.25 | 0.11675 | 0.2 | 0.215
| Std Dev. |  | 0.0089443 | 0.005164 | 0.012816 | 3.0405e-17 | 0.012111
| Rel.Std.Dev. |  | 5.59% | 2.094% | 10.43% | 1.52e-14% | 5.59%
| PDM3 |  | -15.02% | 31.02% | -34.76% | 6.23% | 15.08%
| Z-Score (Absolute) |  | 0.75 | 0.6 | 1.33 | 0.13 | 0.13 | 12.32
| Analytical results for Te in OREAS 353c (Indicative Value 4.98 ppm) | [15 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.11 | 1.12 | 1.13 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 8.96 | < 1 | 1 | < 1 | < 5 | [9 empty] | 1
| 1 | 2 | 8.96 | < 1 |  | < 1 | < 5 | [9 empty] | 5
| 1 | 3 | 8.65 | < 1 |  | < 1 | < 5 | [9 empty] | 16
| 1 | 4 | 9.12 | < 1 |  | < 1 | < 5 | [9 empty] | 4.9833
| 1 | 5 | 9.1 | < 1 |  | < 1 | < 5 | [9 empty] | 23
| 1 | 6 | 9.01 | < 1 | 1 | < 1 | < 5
| Mean |  | 8.9667 |  | 1
| Median |  | 8.985 |  | 1
| Std Dev. |  | 0.16943 |  | 0
| Rel.Std.Dev. |  | 1.89% |  | 0%
| PDM3 |  | 79.93% |  | -79.93%
| Z-Score (Absolute) |  | 10.74 | 0.67 | 0 | 0.67 | 2.02
| Analytical results for Th in OREAS 353c (Indicative Value 0.41 ppm) | [15 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | < 1 | 0.6 | 0.4 | 0.5 | 0.45587 | 0.4 | < 50 | [7 empty] | 1
| 1 | 2 | < 1 | 0.5 | 0.6 | 0.5 | 0.34658 | 0.4 | < 50 | [7 empty] | 18
| 1 | 3 | < 1 | 0.5 | 0.4 | 0.4 | 0.46085 | 0.4 | < 50 | [7 empty] | 16
| 1 | 4 | < 1 | 0.6 | 0.4 | 0.4 | 0.38337 | 0.5 | < 50 | [7 empty] | 0.41403
| 1 | 5 | < 1 | 0.5 | 0.4 | 0.4 | 0.42218 | 0.4 | < 50 | [7 empty] | 24
| 1 | 6 | < 1 | 0.6 | 0.4 | 0.4 | 0.36779 | 0.4 | < 50
| Mean |  |  | 0.55 | 0.43333 | 0.43333 | 0.40611 | 0.41667
| Median |  |  | 0.55 | 0.4 | 0.4 | 0.40277 | 0.4
| Std Dev. |  |  | 0.054772 | 0.08165 | 0.05164 | 0.047455 | 0.040825
| Rel.Std.Dev. |  |  | 9.959% | 18.84% | 11.92% | 11.69% | 9.798%
| PDM3 |  |  | 32.84% | 4.663% | 4.663% | -1.913% | 0.6377%
| Z-Score (Absolute) |  | 1.65 | 2.89 | 0 | 0 | 0.67 | 0.41 | 608.4
| Analytical results for Ti in OREAS 353c (Indicative Value 0.008 wt.%) | [15 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 |  |  | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS |  |  | 3
3
| 1 | 1 | < 0.05 | < 0.01 | < 0.012 | < 0.012 | 0.012 | 0.012 | < 0.012 | < 0.01 | 0.08 | < 0.01 | < 0.01 | 0.009 |  |  | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.012 | < 0.012 | < 0.012 | 0.012 | < 0.012 | < 0.01 | 0.07 | < 0.01 | < 0.01 | 0.0079 |  |  | 6
| 1 | 3 | < 0.05 | < 0.01 | < 0.012 | < 0.012 | 0.012 | 0.012 | < 0.012 | 0.01 | 0.07 | < 0.01 | < 0.01 | 0.0106 |  |  | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.012 | < 0.012 | 0.012 | 0.012 | < 0.012 | < 0.01 | 0.12 | < 0.01 | < 0.01 | 0.0072 |  |  | 0.0082587
| 1 | 5 | < 0.05 | < 0.01 | < 0.012 | < 0.012 | < 0.012 | 0.012 | < 0.012 | < 0.01 | 0.08 | < 0.01 | < 0.01 | 0.008 |  |  | 25
| 1 | 6 | < 0.05 | < 0.01 | < 0.012 | 0.09 | < 0.012 | 0.012 | < 0.012 | < 0.01 | 0.06 | < 0.01 | < 0.01 | 0.0082
| Mean |  |  |  |  | 0.09 | 0.012 | 0.012 |  | 0.01 | 0.08 |  |  | 0.0084833
| Median |  |  |  |  | 0.09 | 0.012 | 0.012 |  | 0.01 | 0.075 |  |  | 0.0081
| Std Dev. | [5 empty] | 2.1246e-18 | 1.9003e-18 |  |  | 0.020976 |  |  | 0.0011873
| Rel.Std.Dev. | [5 empty] | 1.77e-14% | 1.584e-14% |  |  | 26.22% |  |  | 14%
| PDM3 |  |  |  |  | 989.8% | 45.3% | 45.3% |  | 21.08% | 868.7% |  |  | 2.72%
| Z-Score (Absolute) |  | 5.35 | 0.67 | 0.37 | 3.84 | 0.53 | 1.43 | 0.37 | 0.42 | 21.91 | 0.67 | 0.67 | 0.37
| Analytical results for Tl in OREAS 353c (Certified Value 2.02 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.10 | 1.11 | 1.12 | 1.13 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 1.65 | 2.2 | 2.1 | 2 | 2.0043 | 1.9 | [8 empty] | 1
| 1 | 2 | 1.64 | 2.2 | 2.1 | 1.9 | 1.972 | 1.9 | [8 empty] | #N/A
| 1 | 3 | 1.53 | 2.2 | 2.1 | 2 | 2.014 | 1.9 | [8 empty] | 16
| 1 | 4 | 1.61 | 2.2 | 1.9 | 2 | 2.0498 | 2.4 | [8 empty] | 2.0221
| 1 | 5 | 1.62 | 2.1 | 2.1 | 1.9 | 1.9898 | 1.9 | [8 empty] | 67
| 1 | 6 | 1.58 | 2.1 | 2.1 | 2 | 2.0325 | 1.9
| Mean |  | 1.605 | 2.1667 | 2.0667 | 1.9667 | 2.0104 | 1.9833
| Median |  | 1.615 | 2.2 | 2.1 | 2 | 2.0091 | 1.9
| Std Dev. |  | 0.044159 | 0.05164 | 0.08165 | 0.05164 | 0.028266 | 0.20412
| Rel.Std.Dev. |  | 2.751% | 2.383% | 3.951% | 2.626% | 1.406% | 10.29%
| PDM3 |  | -20.63% | 7.15% | 2.205% | -2.74% | -0.5778% | -1.916%
| Z-Score (Absolute) |  | 5.28 | 2.29 | 0.94 | 0.41 | 0.18 | 0.18
| Analytical results for Tm in OREAS 353c (Indicative Value 0.076 ppm) | [15 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.13 | 1.14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | [9 empty] | 3
3
| 1 | 1 | 0.05 | 0.09 | 0.0672 | 0.1 | < 2 | [9 empty] | 1
| 1 | 2 | 0.05 | 0.1 | 0.0645 | 0.1 | < 2 | [9 empty] | 7
| 1 | 3 | 0.05 | 0.1 | 0.0689 | 0.1 | < 2 | [9 empty] | 16
| 1 | 4 | 0.05 | 0.09 | 0.0614 | 0.08 | < 2 | [9 empty] | 0.076133
| 1 | 5 | 0.05 | 0.09 | 0.0655 | 0.1 | < 2 | [9 empty] | 26
| 1 | 6 | 0.05 | 0.1 | 0.0597 | 0.09 | < 2
| Mean |  | 0.05 | 0.095 | 0.064533 | 0.095
| Median |  | 0.05 | 0.095 | 0.065 | 0.1
| Std Dev. |  | 7.6012e-18 | 0.0054772 | 0.0034726 | 0.0083666
| Rel.Std.Dev. |  | 1.52e-14% | 5.766% | 5.381% | 8.807%
| PDM3 |  | -34.33% | 24.78% | -15.24% | 24.78%
| Z-Score (Absolute) |  | 1 | 0 | 0.67 | 0 | 20.03
| Analytical results for U in OREAS 353c (Certified Value 1.26 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 1.19 | 1.3 | 1.3 | 1.1 | 1.1428 | 1.4 | < 50 | [7 empty] | 1
| 1 | 2 | 1.18 | 1.2 | 1.8 | 1 | 1.0804 | 1.3 | < 50 | [7 empty] | #N/A
| 1 | 3 | 1.17 | 1.3 | 1.3 | 1.1 | 1.1075 | 1.4 | < 50 | [7 empty] | 16
| 1 | 4 | 1.23 | 1.3 | 1.7 | 1.1 | 1.2096 | 1.5 | < 50 | [7 empty] | 1.2606
| 1 | 5 | 1.21 | 1.4 | 1.5 | 1.1 | 1.187 | 1.3 | < 50 | [7 empty] | 68
| 1 | 6 | 1.22 | 1.3 | 1.5 | 1.1 | 1.0942 | 1.4 | < 50
| Mean |  | 1.2 | 1.3 | 1.5167 | 1.0833 | 1.1369 | 1.3833
| Median |  | 1.2 | 1.3 | 1.5 | 1.1 | 1.1251 | 1.4
| Std Dev. |  | 0.023664 | 0.063246 | 0.20412 | 0.040825 | 0.052361 | 0.075277
| Rel.Std.Dev. |  | 1.972% | 4.865% | 13.46% | 3.768% | 4.606% | 5.442%
| PDM3 |  | -4.807% | 3.126% | 20.31% | -14.06% | -9.813% | 9.737%
| Z-Score (Absolute) |  | 0.41 | 0 | 0.9 | 0.9 | 0.67 | 0.34 | 97.98
| Analytical results for V in OREAS 353c (Indicative Value < 10 ppm) | [15 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.12 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | [8 empty] | 1
1
| 1 | 1 | < 10 | < 20 | < 10 | < 20 | < 10 | 126 | [8 empty] | 1
| 1 | 2 | < 10 | < 20 | < 10 | < 20 | < 10 | 32 | [8 empty] | 21
| 1 | 3 | < 10 | < 20 | < 10 | < 20 | < 10 | 258 | [8 empty] | 16
| 1 | 4 | < 10 | < 20 | < 10 | < 20 | < 10 | < 5 | [8 empty] | < 10
| 1 | 5 | < 10 | < 20 | < 10 | < 20 | < 10 | < 5 | [8 empty] | 27
| 1 | 6 | < 10 | < 20 | < 10 | < 20 | < 10 | 99
| Mean | [6 empty] | 128.75
| Median | [6 empty] | 112.5
| Std Dev. | [6 empty] | 94.7958
| Rel.Std.Dev. | [6 empty] | 73.63%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 21.35
| Analytical results for W in OREAS 353c (Indicative Value 4.99 ppm) | [15 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 12 | 40000 | 75.64 | < 50 | 3 | 3 | 1.8261 | < 5 | < 20 | [5 empty] | 1
| 1 | 2 | 12 | 50000 | 76.79 | < 50 | 4 | 3 | 1.7583 | < 5 | < 20 | [5 empty] | 8
| 1 | 3 | 12 | 50000 | 77.87 | < 50 | 4 | 2 | 1.884 | < 5 | < 20 | [5 empty] | 16
| 1 | 4 | 12 | 50000 | 77.2 | < 50 | 2 | 5 | 1.9033 | < 5 | < 20 | [5 empty] | 4.9937
| 1 | 5 | 12 | 40000 | 76.79 | < 50 | 4 | 2 | 1.7605 | < 5 | < 20 | [5 empty] | 28
| 1 | 6 | 13 | 50000 | 75.14 | < 50 | 2 | 2 | 1.7178 | < 5 | < 20
| Mean |  | 12.1667 | 46666.6667 | 76.5717 |  | 3.1667 | 2.8333 | 1.8083
| Median |  | 12 | 50000 | 76.79 |  | 3.5 | 2.5 | 1.7933
| Std Dev. |  | 0.40825 | 5163.9778 | 1.0093 |  | 0.98319 | 1.169 | 0.074887
| Rel.Std.Dev. |  | 3.355% | 11.07% | 1.318% |  | 31.05% | 41.26% | 4.141%
| PDM3 |  | 143.6% | 934402% | 1433% |  | -36.59% | -43.26% | -63.79%
| Z-Score (Absolute) |  | 0.19 | 4194.8 | 5.99 | 1.35 | 0.61 | 0.64 | 0.74 | 0.67 | 0
| Analytical results for Y in OREAS 353c (Certified Value 7.79 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 7.92 | 7 | 5.6414 | 10 | 8.5 | 7 | 8.2995 | 7.5 | 15 | [5 empty] | 1
| 1 | 2 | 7.66 | 7.1 | 5.5919 | < 5 | 8.2 | 7.1 | 7.4966 | 7.5 | 9 | [5 empty] | 9
| 1 | 3 | 8.15 | 7.3 | 5.4725 | < 5 | 8.3 | 6.9 | 7.9319 | 7.3 | 11 | [5 empty] | 16
| 1 | 4 | 7.72 | 7.2 | 5.6225 | < 5 | 8.3 | 6.9 | 8.3238 | 7.3 | 2 | [5 empty] | 7.7886
| 1 | 5 | 7.71 | 7.1 | 5.4131 | < 5 | 8.4 | 7 | 7.9219 | 7.2 | 4 | [5 empty] | 69
| 1 | 6 | 7.5 | 7.4 | 5.5516 | < 5 | 8.6 | 6.9 | 7.5562 | 7.1 | 12
| Mean |  | 7.7767 | 7.1833 | 5.5488 | 10 | 8.3833 | 6.9667 | 7.9217 | 7.3167 | 8.8333
| Median |  | 7.715 | 7.15 | 5.5717 | 10 | 8.35 | 6.95 | 7.9269 | 7.3 | 10
| Std Dev. |  | 0.22704 | 0.1472 | 0.089563 |  | 0.1472 | 0.08165 | 0.3518 | 0.16021 | 4.9565
| Rel.Std.Dev. |  | 2.919% | 2.049% | 1.614% |  | 1.756% | 1.172% | 4.441% | 2.19% | 56.11%
| PDM3 |  | -0.1529% | -7.771% | -28.76% | 28.39% | 7.636% | -10.55% | 1.709% | -6.059% | 13.41%
| Z-Score (Absolute) |  | 0.51 | 0.15 | 1.97 | 3.98 | 1.19 | 0.39 | 0.67 | 0 | 1.69
| Analytical results for Yb in OREAS 353c (Certified Value 0.5 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.08 | 1.11 | 1.13 | 1.14 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 0.44 | 0.6 | 0.4735 | 0.6 | 0.5 | < 2 | [8 empty] | 1
| 1 | 2 | 0.43 | 0.6 | 0.4237 | 0.5 | 0.6 | < 2 | [8 empty] | 10
| 1 | 3 | 0.43 | 0.5 | 0.4607 | 0.5 | 0.5 | < 2 | [8 empty] | 16
| 1 | 4 | 0.44 | 0.6 | 0.4574 | 0.6 | 0.5 | < 2 | [8 empty] | 0.50325
| 1 | 5 | 0.45 | 0.6 | 0.4157 | 0.4 | 0.6 | < 2 | [8 empty] | 70
| 1 | 6 | 0.44 | 0.6 | 0.4366 | 0.5 | 0.5 | < 2
| Mean |  | 0.43833 | 0.58333 | 0.4446 | 0.51667 | 0.53333
| Median |  | 0.44 | 0.6 | 0.447 | 0.5 | 0.5
| Std Dev. |  | 0.0075277 | 0.040825 | 0.022778 | 0.075277 | 0.05164
| Rel.Std.Dev. |  | 1.717% | 6.999% | 5.123% | 14.57% | 9.682%
| PDM3 |  | -12.9% | 15.91% | -11.65% | 2.665% | 5.977%
| Z-Score (Absolute) |  | 0.84 | 0.57 | 0.78 | 0.08 | 0.08 | 4.62
| Analytical results for Zn in OREAS 353c (Certified Value 2.86 wt.%) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 2
3
| 1 | 1 | 3.0034 | 2.79 | 2.88 | 2.99 | 2.81 | 2.86 | 2.92 | 2.6821 | 2.5593 | 2.8977 | 2.856 | 2.9557 | 2.89 | 2.718 | 1
| 1 | 2 | 2.9274 | 2.81 | 2.87 | 2.98 | 2.71 | 2.87 | 2.9 | 2.715 | 2.5899 | 3.0043 | 2.72 | 2.8693 | 2.89 | 2.7566 | #N/A
| 1 | 3 | 2.9421 | 2.8 | 2.87 | 2.94 | 2.91 | 2.88 | 2.94 | 2.6912 | 2.541 | 2.9186 | 2.804 | 2.988 | 2.94 | 2.7641 | 16
| 1 | 4 | 2.8968 | 2.77 | 2.89 | 2.96 | 2.91 | 2.88 | 2.8 | 2.658 | 2.5081 | 2.9748 | 2.7301 | 2.9382 | 2.9 | 2.6666 | 2.8594
| 1 | 5 | 2.9688 | 2.75 | 2.92 | 2.99 | 2.86 | 2.87 | 2.89 | 2.6629 | 2.5517 | 2.9269 | 2.7868 | 2.9744 | 2.96 | 2.7784 | 71
| 1 | 6 | 2.8929 | 2.79 | 2.93 | 2.94 | 2.85 | 2.85 | 2.91 | 2.6802 | 2.5971 | 2.993 | 2.8007 | 2.845 | 2.93 | 2.6428
| Mean |  | 2.9386 | 2.785 | 2.8933 | 2.9667 | 2.8417 | 2.8683 | 2.8933 | 2.6816 | 2.5579 | 2.9526 | 2.7829 | 2.9284 | 2.9183 | 2.7211
| Median |  | 2.9348 | 2.79 | 2.885 | 2.97 | 2.855 | 2.87 | 2.905 | 2.6812 | 2.5555 | 2.9508 | 2.7938 | 2.9469 | 2.915 | 2.7373
| Std Dev. |  | 0.042625 | 0.021679 | 0.02582 | 0.023381 | 0.074944 | 0.01169 | 0.048854 | 0.020568 | 0.032762 | 0.043883 | 0.050725 | 0.058228 | 0.029269 | 0.055684
| Rel.Std.Dev. |  | 1.451% | 0.7784% | 0.8924% | 0.7881% | 2.637% | 0.4076% | 1.688% | 0.767% | 1.281% | 1.486% | 1.823% | 1.988% | 1.003% | 2.046%
| PDM3 |  | 2.77% | -2.601% | 1.188% | 3.752% | -0.6191% | 0.3135% | 1.188% | -6.217% | -10.54% | 3.259% | -2.673% | 2.415% | 2.062% | -4.836%
| Z-Score (Absolute) |  | 0.6 | 1 | 0.13 | 0.89 | 0.41 | 0.13 | 0.13 | 2.08 | 3.36 | 0.75 | 1.02 | 0.5 | 0.39 | 1.66
| Analytical results for Zr in OREAS 353c (Indicative Value 4.83 ppm) | [15 empty] | Indicative
Zr | Round | Replicate | Lab | [13 empty] | 1
| No | No | 1.14 | [13 empty] | ppm
|  |  | PF*MS | [13 empty] | 2
2
| 1 | 1 | 6 | [13 empty] | 1
| 1 | 2 | 4 | [13 empty] | 11
| 1 | 3 | 6 | [13 empty] | 16
| 1 | 4 | 4 | [13 empty] | 4.8333
| 1 | 5 | 4 | [13 empty] | 29
| 1 | 6 | 5
| Mean |  | 4.8333
| Median |  | 4.5
| Std Dev. |  | 0.98319
| Rel.Std.Dev. |  | 20.34%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C722: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C758: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C776: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C812: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C830: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C848: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C866: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C884: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C902: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C920: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C938: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C956: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C974: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C992: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1010: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1100: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 353c (Certified Value 0.387 wt.%) | [16 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.42 | 0.4 | 0.35 | 0.39 | 0.38 | 0.38 | 0.39 | 0.47615 | 0.385 | 0.37 | 0.39 | 0.40852 | 0.38 | 0.37 | 0.42 | 1
| 1 | 2 | 0.41 | 0.41 | 0.35 | 0.4 | 0.38 | 0.38 | 0.38 | 0.49039 | 0.382 | 0.38 | 0.34 | 0.4012 | 0.39 | 0.35 | 0.388 | #N/A
| 1 | 3 | 0.42 | 0.41 | 0.35 | 0.39 | 0.38 | 0.38 | 0.39 | 0.4935 | 0.379 | 0.39 | 0.37 | 0.4128 | 0.39 | 0.38 | 0.396 | 16
| 1 | 4 | 0.42 | 0.41 | 0.36 | 0.38 | 0.38 | 0.38 | 0.38 | 0.51175 | 0.379 | 0.38 | 0.37 | 0.40429 | 0.39 | 0.36 | 0.398 | 0.38655
| 1 | 5 | 0.42 | 0.4 | 0.36 | 0.39 | 0.39 | 0.38 | 0.38 | 0.48728 | 0.383 | 0.39 | 0.35 | 0.40336 | 0.39 | 0.37 | 0.396 | 73
| 1 | 6 | 0.43 | 0.41 | 0.39 | 0.4 | 0.39 | 0.38 | 0.38 | 0.48505 | 0.38 | 0.39 | 0.35 | 0.41298 | 0.39 | 0.36 | 0.401
| Mean |  | 0.42 | 0.40667 | 0.36 | 0.39167 | 0.38333 | 0.38 | 0.38333 | 0.49069 | 0.38133 | 0.38333 | 0.36167 | 0.40719 | 0.38833 | 0.365 | 0.39983
| Median |  | 0.42 | 0.41 | 0.355 | 0.39 | 0.38 | 0.38 | 0.38 | 0.48883 | 0.381 | 0.385 | 0.36 | 0.4064 | 0.39 | 0.365 | 0.397
| Std Dev. |  | 0.0063246 | 0.005164 | 0.015492 | 0.0075277 | 0.005164 | 6.0809e-17 | 0.005164 | 0.011886 | 0.0024221 | 0.008165 | 0.018348 | 0.0050151 | 0.0040825 | 0.010488 | 0.010778
| Rel.Std.Dev. |  | 1.506% | 1.27% | 4.303% | 1.922% | 1.347% | 1.6e-14% | 1.347% | 2.422% | 0.6352% | 2.13% | 5.073% | 1.232% | 1.051% | 2.873% | 2.696%
| PDM3 |  | 8.654% | 5.205% | -6.868% | 1.325% | -0.8312% | -1.694% | -0.8312% | 26.94% | -1.349% | -0.8312% | -6.436% | 5.341% | 0.4623% | -5.574% | 3.437%
| Z-Score (Absolute) |  | 1.5 | 0.96 | 0.96 | 0.34 | 0 | 0.14 | 0 | 4.4 | 0.08 | 0 | 0.89 | 0.98 | 0.2 | 0.75 | 0.67
| Analytical results for S in OREAS 353c (Certified Value 14.03 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 |  |  | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC |  |  | 2
3
| 1 | 1 | 13.99 | 14.4 | 13.4 | 14.4 | 13.2 | 13 | 13.9 | 13.8 | 14.32 | 14.5832 | 14.3 | 14.15 | 14.33 |  |  | 1
| 1 | 2 | 14.01 | 14.3 | 13.6 | 14.5 | 13.35 | 13.15 | 13.9 | 14.1 | 14.22 | 14.4936 | 14.5 | 13.94 | 14.53 |  |  | #N/A
| 1 | 3 | 14.01 | 14.4 | 13.55 | 14.55 | 13.4 | 13.55 | 14.1 | 13.9 | 14.17 | 14.5757 | 14.3 | 14.31 | 14.42 |  |  | 16
| 1 | 4 | 14.03 | 14.3 | 13.45 | 14.15 | 13.4 | 12.8 | 13.85 | 14 | 14.26 | 14.5451 | 14.2 | 13.94 | 14.62 |  |  | 14.0336
| 1 | 5 | 14.01 | 14.4 | 13.75 | 14.85 | 13.45 | 12.85 | 13.8 | 13.9 | 14.38 | 14.5036 | 14.4 | 14.24 | 14.67 |  |  | 74
| 1 | 6 | 14.01 | 14.4 | 13.45 | 14.6 | 13.5 | 13.35 | 13.7 | 13.7 | 14.28 | 14.6078 | 14.3 | 13.94 | 14.43
| Mean |  | 14.01 | 14.3667 | 13.5333 | 14.5083 | 13.3833 | 13.1167 | 13.875 | 13.9 | 14.2717 | 14.5515 | 14.3333 | 14.0867 | 14.5
| Median |  | 14.01 | 14.4 | 13.5 | 14.525 | 13.4 | 13.075 | 13.875 | 13.9 | 14.27 | 14.5604 | 14.3 | 14.045 | 14.48
| Std Dev. |  | 0.012649 | 0.05164 | 0.1291 | 0.23112 | 0.10328 | 0.29269 | 0.13323 | 0.14142 | 0.073869 | 0.045706 | 0.10328 | 0.16848 | 0.12992
| Rel.Std.Dev. |  | 0.09029% | 0.3594% | 0.9539% | 1.593% | 0.7717% | 2.231% | 0.9602% | 1.017% | 0.5176% | 0.3141% | 0.7206% | 1.196% | 0.896%
| PDM3 |  | -0.168% | 2.374% | -3.565% | 3.383% | -4.633% | -6.534% | -1.13% | -0.9518% | 1.697% | 3.69% | 2.136% | 0.3783% | 3.324%
| Z-Score (Absolute) |  | 0.18 | 0.67 | 1.33 | 1.02 | 1.69 | 2.34 | 0.51 | 0.45 | 0.45 | 1.12 | 0.59 | 0 | 1
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.