## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 284
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.40389 | 0.39872 | 0.40906 | 0.40148 | 0.40629
PhotonAssay (recommended gross mass* 330±15 g)
Au, Gold (ppm) | 0.4109 | 0.40313 | 0.41867 | 0.41018 | 0.41162
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 0.39429 | 0.38613 | 0.40246 | 0.39172 | 0.39686
Cyanide Leach
Au, Gold (ppm) | 0.38534 | 0.36908 | 0.4016 | 0.38445 | 0.38622
4-Acid Digestion
Ag, Silver (ppm) | 0.095207 | 0.079551 | 0.11086 | IND | IND
Al, Aluminium (wt.%) | 7.7416 | 7.5002 | 7.9829 | 7.5993 | 7.8839
As, Arsenic (ppm) | 179.9773 | 172.5367 | 187.4178 | 173.6595 | 186.295
Ba, Barium (ppm) | 1044.5144 | 1004.8128 | 1084.2159 | 1023.1237 | 1065.905
Be, Beryllium (ppm) | 2.56 | 2.4164 | 2.7036 | 2.4546 | 2.6654
Bi, Bismuth (ppm) | 0.66682 | 0.53797 | 0.79568 | 0.57301 | 0.76063
Ca, Calcium (wt.%) | 1.8676 | 1.8074 | 1.9277 | 1.8221 | 1.913
Cd, Cadmium (ppm) | 0.2027 | 0.17053 | 0.23487 | 0.17941 | 0.226
Ce, Cerium (ppm) | 78.9981 | 73.9528 | 84.0434 | 76.4551 | 81.541
Co, Cobalt (ppm) | 12.2823 | 11.5686 | 12.9961 | 11.8615 | 12.7032
Cr, Chromium (ppm) | 47.7435 | 44.2359 | 51.2512 | 45.8654 | 49.6217
Cs, Caesium (ppm) | 11.3727 | 10.7865 | 11.9589 | 11.0384 | 11.707
Cu, Copper (ppm) | 27.5974 | 25.6917 | 29.5031 | 26.5673 | 28.6274
Dy, Dysprosium (ppm) | 3.9404 | 3.657 | 4.2238 | 3.7219 | 4.1589
Er, Erbium (ppm) | 1.4199 | 1.2422 | 1.5977 | 1.3569 | 1.4829
Eu, Europium (ppm) | 1.4774 | 1.3148 | 1.6401 | 1.3772 | 1.5777
Fe, Iron (wt.%) | 3.1542 | 3.0507 | 3.2577 | 3.081 | 3.2275
Ga, Gallium (ppm) | 21.0112 | 19.7651 | 22.2574 | 20.2536 | 21.7689
Gd, Gadolinium (ppm) | 6.1892 | 5.5388 | 6.8396 | 5.9402 | 6.4381
Hf, Hafnium (ppm) | 1.9285 | 1.763 | 2.094 | 1.8074 | 2.0496
Ho, Holmium (ppm) | 0.62789 | 0.577 | 0.67879 | 0.58955 | 0.66624
In, Indium (ppm) | 0.069018 | 0.062632 | 0.075405 | 0.063528 | 0.074509
K, Potassium (wt.%) | 2.9261 | 2.8258 | 3.0263 | 2.8593 | 2.9928
La, Lanthanum (ppm) | 37.9609 | 35.7292 | 40.1926 | 36.6832 | 39.2386
Li, Lithium (ppm) | 53.169 | 50.4891 | 55.8489 | 51.805 | 54.533
Lu, Lutetium (ppm) | 0.16179 | 0.13938 | 0.18419 | IND | IND
Mg, Magnesium (wt.%) | 0.85244 | 0.82471 | 0.88017 | 0.83495 | 0.86992
Mn, Manganese (wt.%) | 0.037781 | 0.036311 | 0.03925 | 0.03688 | 0.038681
Mo, Molybdenum (ppm) | 2.5769 | 2.2949 | 2.859 | 2.3546 | 2.7992
Na, Sodium (wt.%) | 2.0414 | 1.9824 | 2.1005 | 1.9977 | 2.0852
Nb, Niobium (ppm) | 12.8932 | 12.2957 | 13.4907 | 12.5073 | 13.2791
Nd, Neodymium (ppm) | 35.6018 | 32.7145 | 38.4892 | 33.9928 | 37.2108
Ni, Nickel (ppm) | 18.8507 | 17.8406 | 19.8608 | 18.3317 | 19.3698
P, Phosphorus (wt.%) | 0.084071 | 0.080949 | 0.087193 | 0.081967 | 0.086175
Pb, Lead (ppm) | 23.5631 | 22.4076 | 24.7187 | 22.6085 | 24.5178
Pr, Praseodymium (ppm) | 9.2837 | 8.3016 | 10.2658 | 8.7958 | 9.7716
Rb, Rubidium (ppm) | 169.6048 | 161.0785 | 178.131 | 164.7058 | 174.5037
S, Sulphur (wt.%) | 0.26173 | 0.25038 | 0.27308 | 0.25417 | 0.26929
Sb, Antimony (ppm) | 0.46625 | 0.42497 | 0.50754 | 0.43718 | 0.49533
Sc, Scandium (ppm) | 9.0308 | 8.5072 | 9.5543 | 8.7575 | 9.3041
Sm, Samarium (ppm) | 7.2704 | 6.2388 | 8.302 | 6.9624 | 7.5785
Sn, Tin (ppm) | 4.7877 | 4.5138 | 5.0616 | 4.5751 | 5.0002
Sr, Strontium (ppm) | 161.3172 | 155.179 | 167.4555 | 158.0302 | 164.6043
Ta, Tantalum (ppm) | 1.09 | 1.0147 | 1.1654 | 1.0548 | 1.1253
Tb, Terbium (ppm) | 0.79188 | 0.71312 | 0.87065 | 0.75307 | 0.83069
Te, Tellurium (ppm) | 0.056015 | 0.013038 | 0.098992 | IND | IND
Th, Thorium (ppm) | 15.2733 | 14.676 | 15.8706 | 14.7357 | 15.8109
Ti, Titanium (wt.%) | 0.37309 | 0.35998 | 0.38621 | 0.36242 | 0.38377
Tl, Thallium (ppm) | 0.95386 | 0.91153 | 0.9962 | 0.92351 | 0.98422
Tm, Thulium (ppm) | 0.18907 | 0.16777 | 0.21036 | IND | IND
U, Uranium (ppm) | 4.0929 | 3.5274 | 4.6584 | 3.8184 | 4.3675
V, Vanadium (ppm) | 62.4928 | 60.0207 | 64.9648 | 60.6989 | 64.2867
W, Tungsten (ppm) | 16.0541 | 14.9408 | 17.1674 | 15.0808 | 17.0274
Y, Yttrium (ppm) | 15.7883 | 14.7986 | 16.7779 | 15.2259 | 16.3506
Yb, Ytterbium (ppm) | 1.1739 | 1.0186 | 1.3292 | 1.0695 | 1.2783
Zn, Zinc (ppm) | 82.0751 | 79.0639 | 85.0862 | 79.9752 | 84.1749
Zr, Zirconium (ppm) | 58.5493 | 54.2375 | 62.8612 | 55.5901 | 61.5086
Aqua Regia Digestion
Ag, Silver (ppm) | 0.082347 | 0.069695 | 0.094999 | 0.073399 | 0.091296
Al, Aluminium (wt.%) | 2.1645 | 2.0753 | 2.2536 | 2.1048 | 2.2242
As, Arsenic (ppm) | 174.1327 | 167.2903 | 180.9751 | 170.0762 | 178.1892
Ba, Barium (ppm) | 510.0781 | 486.6471 | 533.509 | 495.7562 | 524.3999
Be, Beryllium (ppm) | 1.6593 | 1.5526 | 1.766 | 1.595 | 1.7235
Bi, Bismuth (ppm) | 0.69424 | 0.52594 | 0.86254 | 0.58375 | 0.80473
Ca, Calcium (wt.%) | 0.57184 | 0.55085 | 0.59283 | 0.55476 | 0.58892
Cd, Cadmium (ppm) | 0.083833 | 0.068324 | 0.099342 | IND | IND
Ce, Cerium (ppm) | 40.4456 | 36.482 | 44.4091 | 38.8242 | 42.067
Co, Cobalt (ppm) | 11.914 | 11.317 | 12.511 | 11.5549 | 12.273
Cr, Chromium (ppm) | 52.1665 | 49.7074 | 54.6255 | 50.7636 | 53.5693
Cs, Caesium (ppm) | 9.4776 | 9.1084 | 9.8467 | 9.1326 | 9.8226
Cu, Copper (ppm) | 27.4121 | 26.1187 | 28.7055 | 26.5093 | 28.3149
Dy, Dysprosium (ppm) | 2.8114 | 2.4003 | 3.2225 | 2.6225 | 3.0002
Er, Erbium (ppm) | 0.99932 | 0.82864 | 1.17 | IND | IND
Eu, Europium (ppm) | 0.38513 | 0.27686 | 0.4934 | 0.32886 | 0.44139
Fe, Iron (wt.%) | 3.0825 | 2.9449 | 3.2201 | 3.0075 | 3.1576
Ga, Gallium (ppm) | 9.7618 | 9.285 | 10.2386 | 9.356 | 10.1676
Gd, Gadolinium (ppm) | 4.0522 | 2.8968 | 5.2076 | 3.6566 | 4.4478
Ge, Germanium (ppm) | 0.12194 | 0.086705 | 0.15718 | IND | IND
Hf, Hafnium (ppm) | 0.34175 | 0.30399 | 0.37951 | 0.32625 | 0.35725
Ho, Holmium (ppm) | 0.40563 | 0.35731 | 0.45394 | 0.3886 | 0.42266
In, Indium (ppm) | 0.066433 | 0.056441 | 0.076425 | 0.060339 | 0.072528
K, Potassium (wt.%) | 1.0152 | 0.97685 | 1.0535 | 0.99093 | 1.0394
La, Lanthanum (ppm) | 19.2174 | 17.2317 | 21.203 | 18.2222 | 20.2126
Li, Lithium (ppm) | 44.4363 | 42.5695 | 46.3031 | 43.1966 | 45.6759
Lu, Lutetium (ppm) | 0.095156 | 0.084296 | 0.10602 | IND | IND
Mg, Magnesium (wt.%) | 0.80997 | 0.77799 | 0.84196 | 0.79012 | 0.82983
Mn, Manganese (wt.%) | 0.032725 | 0.031293 | 0.034156 | 0.031785 | 0.033664
Mo, Molybdenum (ppm) | 2.3721 | 2.1524 | 2.5918 | 2.1327 | 2.6115
Na, Sodium (wt.%) | 0.1593 | 0.14582 | 0.17278 | 0.15032 | 0.16828
Nb, Niobium (ppm) | 0.78237 | 0.59277 | 0.97197 | 0.72136 | 0.84338
Ni, Nickel (ppm) | 18.6571 | 17.8 | 19.5141 | 18.0611 | 19.253
P, Phosphorus (wt.%) | 0.064041 | 0.06189 | 0.066191 | 0.062634 | 0.065447
Pb, Lead (ppm) | 4.794 | 4.414 | 5.1741 | 4.5954 | 4.9927
Rb, Rubidium (ppm) | 105.167 | 99.9429 | 110.3912 | 102.9392 | 107.3949
S, Sulphur (wt.%) | 0.255 | 0.24043 | 0.26956 | 0.24677 | 0.26322
Sb, Antimony (ppm) | 0.24621 | 0.21918 | 0.27323 | 0.20641 | 0.28601
Sc, Scandium (ppm) | 8.1571 | 7.8756 | 8.4386 | 7.8666 | 8.4476
Se, Selenium (ppm) | 0.39741 | 0.14812 | 0.64669 | IND | IND
Sn, Tin (ppm) | 3.5449 | 3.3455 | 3.7443 | 3.4193 | 3.6704
Sr, Strontium (ppm) | 19.313 | 17.8926 | 20.7334 | 18.6249 | 20.0011
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.57463 | 0.50827 | 0.64098 | 0.54131 | 0.60794
Te, Tellurium (ppm) | 0.049128 | 0.028492 | 0.069765 | IND | IND
Th, Thorium (ppm) | 8.5151 | 7.5637 | 9.4664 | 8.1191 | 8.911
Ti, Titanium (wt.%) | 0.27265 | 0.25887 | 0.28644 | 0.26281 | 0.28249
Tl, Thallium (ppm) | 0.67857 | 0.64581 | 0.71132 | 0.64972 | 0.70741
U, Uranium (ppm) | 3.4653 | 2.9853 | 3.9452 | 3.1973 | 3.7332
V, Vanadium (ppm) | 60.7629 | 58.7215 | 62.8043 | 59.3217 | 62.2042
W, Tungsten (ppm) | 10.1375 | 9.296 | 10.979 | 9.7422 | 10.5327
Y, Yttrium (ppm) | 10.8395 | 10.2548 | 11.4242 | 10.577 | 11.102
Yb, Ytterbium (ppm) | 0.7082 | 0.6819 | 0.7345 | 0.65515 | 0.76126
Zn, Zinc (ppm) | 77.0612 | 74.1333 | 79.9891 | 75.2029 | 78.9195
Zr, Zirconium (ppm) | 8.274 | 7.7548 | 8.7932 | 7.8986 | 8.6493
Infrared Combustion
S, Sulphur (wt.%) | 0.25716 | 0.24585 | 0.26848 | 0.24633 | 0.268
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 284
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 9.6451 | Hg | ppm | 0.13333 | Se | ppm | 0.70278
Ge | ppm | 0.2682 | Re | ppm | 0.0015
Aqua Regia Digestion
B | ppm | 9.7769 | Pr | ppm | 5.5062 | Sm | ppm | 4.3681
Hg | ppm | 0.013738 | Pt | ppb | < 5 | Tm | ppm | 0.11673
Nd | ppm | 22.4359 | Re | ppm | 0.0011909
Pd | ppb | 19.5 | Si | wt.% | 0.13345
Infrared Combustion
C | wt.% | 0.12917
Borate Fusion XRF
Al2O3 | wt.% | 15.165 | Fe | wt.% | 3.235 | S | wt.% | 0.262
As | ppm | 170 | K2O | wt.% | 3.55 | SiO2 | wt.% | 67.515
Ba | ppm | 1065 | MgO | wt.% | 1.485 | Sn | ppm | < 10
CaO | wt.% | 2.635 | MnO | wt.% | 0.0505 | Sr | ppm | 155
Cl | ppm | 80 | Na2O | wt.% | 2.76 | TiO2 | wt.% | 0.633
Co | ppm | 10 | Ni | ppm | 25 | V | ppm | 60
Cr | ppm | 60 | P | wt.% | 0.086 | Zn | ppm | 80
Cu | ppm | 35 | Pb | ppm | 20 | Zr | ppm | 240
Thermogravimetry
LOI1000 | wt.% | 1.17
Laser Ablation ICP-MS
Ag | ppm | 0.1 | Hf | ppm | 7.055 | Sm | ppm | 7.88
As | ppm | 165 | Ho | ppm | 1.145 | Sn | ppm | 4.5
Ba | ppm | 998.5 | In | ppm | 0.05 | Sr | ppm | 157
Be | ppm | 2.6 | La | ppm | 39.65 | Ta | ppm | 1.105
Bi | ppm | 0.6 | Lu | ppm | 0.39 | Tb | ppm | 1.105
Cd | ppm | 0.1 | Mn | wt.% | 0.03765 | Te | ppm | < 0.2
Ce | ppm | 79.15 | Mo | ppm | 2.4 | Th | ppm | 15.3
Co | ppm | 12.25 | Nb | ppm | 13 | Ti | wt.% | 0.374
Cr | ppm | 51 | Nd | ppm | 37.5 | Tl | ppm | 0.8
Cs | ppm | 11.25 | Ni | ppm | 19 | Tm | ppm | 0.44
Cu | ppm | 28 | Pb | ppm | 23 | U | ppm | 4.695
Dy | ppm | 6.065 | Pr | ppm | 9.755 | V | ppm | 64.3
Er | ppm | 3.12 | Rb | ppm | 164.5 | W | ppm | 17.5
Eu | ppm | 1.455 | Re | ppm | < 0.01 | Y | ppm | 30.6
Ga | ppm | 20 | Sb | ppm | 0.4 | Yb | ppm | 2.86
Gd | ppm | 7.065 | Sc | ppm | 8.3 | Zn | ppm | 75
Ge | ppm | 1.375 | Se | ppm | < 5 | Zr | ppm | 247.5
3-Acid Digestion (no HF)
Ag | ppm | 0.22333 | Gd | ppm | 4.8167 | S | wt.% | 0.2805
Al2O3 | wt.% | 15.85 | Hf | ppm | 2.4333 | Sc | ppm | 8.4
Ba | ppm | 1066.6667 | Ho | ppm | 0.51 | Sm | ppm | 6.6667
Be | ppm | 3 | K2O | wt.% | 3.6867 | Sn | ppm | 4.2133
Bi | ppm | 0.6 | La | ppm | 38.8333 | Sr | ppm | 165
CaO | wt.% | 2.6417 | Li | ppm | 52.6667 | Ta | ppm | 0.93167
Cd | ppm | 0.23333 | MgO | wt.% | 1.5317 | Tb | ppm | 0.65333
Ce | ppm | 77.6667 | MnO | wt.% | 0.048833 | Th | ppm | 15.45
Co | ppm | 12.65 | Mo | ppm | 2.9117 | TiO2 | wt.% | 0.65967
Cr | ppm | 44.3333 | Na2O | wt.% | 2.815 | U | ppm | 3.8117
Cs | ppm | 9.3333 | Nb | ppm | 13.2667 | V | ppm | 66
Cu | ppm | 27.4833 | Nd | ppm | 33.9833 | W | ppm | 18.2167
Dy | ppm | 3.6067 | Ni | ppm | 22.7833 | Y | ppm | 13.7333
Er | ppm | 1.4 | P2O5 | wt.% | 0.20067 | Yb | ppm | 1.0133
Eu | ppm | 1.6033 | Pb | ppm | 23.85 | Zn | ppm | 81.8333
Fe2O3 | wt.% | 4.6783 | Pr | ppm | 8.5667 | Zr | ppm | 55.3333
Ga | ppm | 18.5833 | Rb | ppm | 194
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 284
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.40389 | 0.015602 | 0.37268 | 0.43509 | 0.35708 | 0.45069 | 3.863% | 7.726% | 11.59% | 0.38369 | 0.42408
PhotonAssay
Au, ppm | 0.4109 | 0.03575 | 0.3394 | 0.4824 | 0.30365 | 0.51815 | 8.7% | 17.4% | 26.1% | 0.39036 | 0.43145
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.39429 | 0.019295 | 0.3557 | 0.43288 | 0.33641 | 0.45218 | 4.894% | 9.787% | 14.68% | 0.37458 | 0.41401
Cyanide Leach
Au, ppm | 0.38534 | 0.036548 | 0.31224 | 0.45843 | 0.27569 | 0.49498 | 9.485% | 18.97% | 28.45% | 0.36607 | 0.4046
4-Acid Digestion
Ag, ppm | 0.095207 | 0.012389 | 0.070429 | 0.11998 | 0.05804 | 0.13237 | 13.01% | 26.02% | 39.04% | 0.090446 | 0.099967
Al, wt.% | 7.7416 | 0.28766 | 7.1663 | 8.3169 | 6.8786 | 8.6045 | 3.716% | 7.431% | 11.15% | 7.3545 | 8.1287
As, ppm | 179.9773 | 7.0816 | 165.8141 | 194.1404 | 158.7325 | 201.222 | 3.935% | 7.869% | 11.8% | 170.9784 | 188.9761
Ba, ppm | 1044.5144 | 49.3826 | 945.7493 | 1143.2795 | 896.3667 | 1192.662 | 4.728% | 9.456% | 14.18% | 992.2886 | 1096.7401
Be, ppm | 2.56 | 0.17901 | 2.202 | 2.918 | 2.023 | 3.0971 | 6.993% | 13.99% | 20.98% | 2.432 | 2.688
Bi, ppm | 0.66682 | 0.10167 | 0.46349 | 0.87016 | 0.36182 | 0.97182 | 15.25% | 30.49% | 45.74% | 0.63348 | 0.70016
Ca, wt.% | 1.8676 | 0.080943 | 1.7057 | 2.0294 | 1.6247 | 2.1104 | 4.334% | 8.668% | 13% | 1.7742 | 1.9609
Cd, ppm | 0.2027 | 0.021181 | 0.16034 | 0.24506 | 0.13916 | 0.26625 | 10.45% | 20.9% | 31.35% | 0.19257 | 0.21284
Ce, ppm | 78.9981 | 4.6663 | 69.6655 | 88.3306 | 64.9993 | 92.9969 | 5.907% | 11.81% | 17.72% | 75.0482 | 82.948
Co, ppm | 12.2823 | 0.71146 | 10.8594 | 13.7053 | 10.148 | 14.4167 | 5.793% | 11.59% | 17.38% | 11.6682 | 12.8965
Cr, ppm | 47.7435 | 3.9634 | 39.8167 | 55.6704 | 35.8533 | 59.6338 | 8.301% | 16.6% | 24.9% | 45.3564 | 50.1307
Cs, ppm | 11.3727 | 0.55575 | 10.2612 | 12.4842 | 9.7054 | 13.0399 | 4.887% | 9.773% | 14.66% | 10.804 | 11.9413
Cu, ppm | 27.5974 | 2.3237 | 22.9499 | 32.2448 | 20.6262 | 34.5686 | 8.42% | 16.84% | 25.26% | 26.2175 | 28.9772
Dy, ppm | 3.9404 | 0.19015 | 3.5601 | 4.3207 | 3.37 | 4.5109 | 4.826% | 9.651% | 14.48% | 3.7434 | 4.1374
Er, ppm | 1.4199 | 0.082642 | 1.2546 | 1.5852 | 1.172 | 1.6679 | 5.82% | 11.64% | 17.46% | 1.3489 | 1.4909
Eu, ppm | 1.4774 | 0.087629 | 1.3022 | 1.6527 | 1.2145 | 1.7403 | 5.931% | 11.86% | 17.79% | 1.4035 | 1.5513
Fe, wt.% | 3.1542 | 0.099389 | 2.9554 | 3.353 | 2.856 | 3.4524 | 3.151% | 6.302% | 9.453% | 2.9965 | 3.3119
Ga, ppm | 21.0112 | 1.3957 | 18.2198 | 23.8027 | 16.8241 | 25.1984 | 6.643% | 13.29% | 19.93% | 19.9607 | 22.0618
Gd, ppm | 6.1892 | 0.26013 | 5.6689 | 6.7094 | 5.4088 | 6.9696 | 4.203% | 8.406% | 12.61% | 5.8797 | 6.4986
Hf, ppm | 1.9285 | 0.14242 | 1.6436 | 2.2133 | 1.5012 | 2.3557 | 7.385% | 14.77% | 22.16% | 1.8321 | 2.0249
Ho, ppm | 0.62789 | 0.039013 | 0.54987 | 0.70592 | 0.51085 | 0.74493 | 6.213% | 12.43% | 18.64% | 0.5965 | 0.65929
In, ppm | 0.069018 | 0.0043949 | 0.060229 | 0.077808 | 0.055834 | 0.082203 | 6.368% | 12.74% | 19.1% | 0.065568 | 0.072469
K, wt.% | 2.9261 | 0.10353 | 2.719 | 3.1331 | 2.6155 | 3.2367 | 3.538% | 7.076% | 10.61% | 2.7798 | 3.0724
La, ppm | 37.9609 | 1.6861 | 34.5886 | 41.3332 | 32.9025 | 43.0193 | 4.442% | 8.884% | 13.33% | 36.0628 | 39.8589
Li, ppm | 53.169 | 3.2344 | 46.7001 | 59.6379 | 43.4656 | 62.8723 | 6.083% | 12.17% | 18.25% | 50.5105 | 55.8274
Lu, ppm | 0.16179 | 0.018211 | 0.12536 | 0.19821 | 0.10715 | 0.21642 | 11.26% | 22.51% | 33.77% | 0.1537 | 0.16988
Mg, wt.% | 0.85244 | 0.033549 | 0.78534 | 0.91954 | 0.75179 | 0.95309 | 3.936% | 7.871% | 11.81% | 0.80982 | 0.89506
Mn, wt.% | 0.037781 | 0.0017325 | 0.034316 | 0.041245 | 0.032583 | 0.042978 | 4.586% | 9.171% | 13.76% | 0.035892 | 0.03967
Mo, ppm | 2.5769 | 0.19973 | 2.1775 | 2.9764 | 1.9777 | 3.1761 | 7.751% | 15.5% | 23.25% | 2.4481 | 2.7058
Na, wt.% | 2.0414 | 0.069597 | 1.9023 | 2.1806 | 1.8327 | 2.2502 | 3.409% | 6.818% | 10.23% | 1.9394 | 2.1435
Nb, ppm | 12.8932 | 0.7563 | 11.3806 | 14.4058 | 10.6243 | 15.1621 | 5.866% | 11.73% | 17.6% | 12.2485 | 13.5379
Nd, ppm | 35.6018 | 2.0856 | 31.4306 | 39.7731 | 29.345 | 41.8587 | 5.858% | 11.72% | 17.57% | 33.8217 | 37.3819
Ni, ppm | 18.8507 | 0.87017 | 17.1104 | 20.5911 | 16.2402 | 21.4612 | 4.616% | 9.232% | 13.85% | 17.9082 | 19.7933
P, wt.% | 0.084071 | 0.0029891 | 0.078093 | 0.090049 | 0.075104 | 0.093038 | 3.555% | 7.111% | 10.67% | 0.079868 | 0.088275
Pb, ppm | 23.5631 | 1.6345 | 20.2942 | 26.8321 | 18.6598 | 28.4665 | 6.937% | 13.87% | 20.81% | 22.385 | 24.7413
Pr, ppm | 9.2837 | 0.51534 | 8.253 | 10.3144 | 7.7377 | 10.8297 | 5.551% | 11.1% | 16.65% | 8.8195 | 9.7479
Rb, ppm | 169.6048 | 9.8282 | 149.9483 | 189.2612 | 140.1201 | 199.0894 | 5.795% | 11.59% | 17.38% | 161.1245 | 178.085
S, wt.% | 0.26173 | 0.011134 | 0.23946 | 0.284 | 0.22833 | 0.29513 | 4.254% | 8.508% | 12.76% | 0.24864 | 0.27482
Sb, ppm | 0.46625 | 0.02818 | 0.40989 | 0.52261 | 0.38171 | 0.55079 | 6.044% | 12.09% | 18.13% | 0.44294 | 0.48957
Sc, ppm | 9.0308 | 0.56735 | 7.8961 | 10.1655 | 7.3287 | 10.7328 | 6.282% | 12.56% | 18.85% | 8.5792 | 9.4823
Sm, ppm | 7.2704 | 0.46809 | 6.3343 | 8.2066 | 5.8662 | 8.6747 | 6.438% | 12.88% | 19.31% | 6.9069 | 7.634
Sn, ppm | 4.7877 | 0.23562 | 4.3164 | 5.2589 | 4.0808 | 5.4945 | 4.921% | 9.843% | 14.76% | 4.5483 | 5.0271
Sr, ppm | 161.3172 | 6.8217 | 147.6738 | 174.9606 | 140.8521 | 181.7823 | 4.229% | 8.457% | 12.69% | 153.2514 | 169.3831
Ta, ppm | 1.09 | 0.057863 | 0.9743 | 1.2058 | 0.91644 | 1.2636 | 5.308% | 10.62% | 15.93% | 1.0355 | 1.1445
Tb, ppm | 0.79188 | 0.064987 | 0.66191 | 0.92186 | 0.59692 | 0.98684 | 8.207% | 16.41% | 24.62% | 0.75229 | 0.83148
Te, ppm | 0.056015 | 0.024879 | 0.0062571 | 0.10577 | 0 | 0.13065 | 44.41% | 88.83% | 133.2% | 0.053214 | 0.058816
Th, ppm | 15.2733 | 0.68357 | 13.9062 | 16.6405 | 13.2226 | 17.324 | 4.476% | 8.951% | 13.43% | 14.5097 | 16.037
Ti, wt.% | 0.37309 | 0.011867 | 0.34936 | 0.39683 | 0.33749 | 0.4087 | 3.181% | 6.362% | 9.542% | 0.35444 | 0.39175
Tl, ppm | 0.95386 | 0.05006 | 0.85375 | 1.054 | 0.80369 | 1.104 | 5.248% | 10.5% | 15.74% | 0.90617 | 1.0016
Tm, ppm | 0.18907 | 0.014049 | 0.16097 | 0.21716 | 0.14692 | 0.23121 | 7.431% | 14.86% | 22.29% | 0.17961 | 0.19852
U, ppm | 4.0929 | 0.43696 | 3.219 | 4.9669 | 2.7821 | 5.4038 | 10.68% | 21.35% | 32.03% | 3.8883 | 4.2976
V, ppm | 62.4928 | 2.6385 | 57.2158 | 67.7697 | 54.5773 | 70.4082 | 4.222% | 8.444% | 12.67% | 59.3681 | 65.6174
W, ppm | 16.0541 | 0.96983 | 14.1145 | 17.9938 | 13.1446 | 18.9636 | 6.041% | 12.08% | 18.12% | 15.2514 | 16.8568
Y, ppm | 15.7883 | 1.1292 | 13.53 | 18.0466 | 12.4008 | 19.1757 | 7.152% | 14.3% | 21.46% | 14.9989 | 16.5777
Yb, ppm | 1.1739 | 0.14464 | 0.88462 | 1.4632 | 0.73997 | 1.6078 | 12.32% | 24.64% | 36.96% | 1.1152 | 1.2326
Zn, ppm | 82.0751 | 3.3505 | 75.374 | 88.7761 | 72.0235 | 92.1266 | 4.082% | 8.164% | 12.25% | 77.9713 | 86.1788
Zr, ppm | 58.5493 | 5.4242 | 47.7009 | 69.3977 | 42.2767 | 74.822 | 9.264% | 18.53% | 27.79% | 55.6219 | 61.4768
Aqua Regia Digestion
Ag, ppm | 0.082347 | 0.009264 | 0.063819 | 0.10088 | 0.054555 | 0.11014 | 11.25% | 22.5% | 33.75% | 0.07823 | 0.086465
Al, wt.% | 2.1645 | 0.16205 | 1.8404 | 2.4886 | 1.6783 | 2.6506 | 7.487% | 14.97% | 22.46% | 2.0562 | 2.2727
As, ppm | 174.1327 | 7.8557 | 158.4214 | 189.844 | 150.5657 | 197.6997 | 4.511% | 9.023% | 13.53% | 165.4261 | 182.8393
Ba, ppm | 510.0781 | 33.585 | 442.9081 | 577.248 | 409.3231 | 610.833 | 6.584% | 13.17% | 19.75% | 484.5742 | 535.582
Be, ppm | 1.6593 | 0.094211 | 1.4708 | 1.8477 | 1.3766 | 1.9419 | 5.678% | 11.36% | 17.03% | 1.5763 | 1.7422
Bi, ppm | 0.69424 | 0.084176 | 0.52589 | 0.86259 | 0.44171 | 0.94677 | 12.12% | 24.25% | 36.37% | 0.65953 | 0.72895
Ca, wt.% | 0.57184 | 0.032801 | 0.50624 | 0.63744 | 0.47343 | 0.67024 | 5.736% | 11.47% | 17.21% | 0.54325 | 0.60043
Cd, ppm | 0.083833 | 0.0091399 | 0.065553 | 0.10211 | 0.056414 | 0.11125 | 10.9% | 21.8% | 32.71% | 0.079642 | 0.088025
Ce, ppm | 40.4456 | 4.767 | 30.9115 | 49.9796 | 26.1445 | 54.7467 | 11.79% | 23.57% | 35.36% | 38.4233 | 42.4679
Co, ppm | 11.914 | 0.61686 | 10.6803 | 13.1477 | 10.0634 | 13.7645 | 5.178% | 10.36% | 15.53% | 11.3183 | 12.5097
Cr, ppm | 52.1665 | 2.3939 | 47.3787 | 56.9543 | 44.9848 | 59.3482 | 4.589% | 9.178% | 13.77% | 49.5581 | 54.7748
Cs, ppm | 9.4776 | 0.29219 | 8.8932 | 10.0619 | 8.601 | 10.3541 | 3.083% | 6.166% | 9.249% | 9.0037 | 9.9514
Cu, ppm | 27.4121 | 1.5946 | 24.2229 | 30.6013 | 22.6283 | 32.1959 | 5.817% | 11.63% | 17.45% | 26.0415 | 28.7827
Dy, ppm | 2.8114 | 0.27042 | 2.2705 | 3.3522 | 2.0001 | 3.6226 | 9.619% | 19.24% | 28.86% | 2.6708 | 2.9519
Er, ppm | 0.99932 | 0.1123 | 0.77472 | 1.2239 | 0.66243 | 1.3362 | 11.24% | 22.47% | 33.71% | 0.94935 | 1.0493
Eu, ppm | 0.38513 | 0.075589 | 0.23395 | 0.5363 | 0.15836 | 0.61189 | 19.63% | 39.25% | 58.88% | 0.36587 | 0.40438
Fe, wt.% | 3.0825 | 0.16004 | 2.7624 | 3.4026 | 2.6024 | 3.5626 | 5.192% | 10.38% | 15.58% | 2.9284 | 3.2366
Ga, ppm | 9.7618 | 0.58836 | 8.5851 | 10.9385 | 7.9967 | 11.5268 | 6.027% | 12.05% | 18.08% | 9.2737 | 10.2499
Gd, ppm | 4.0522 | 0.81507 | 2.4221 | 5.6824 | 1.607 | 6.4974 | 20.11% | 40.23% | 60.34% | 3.8496 | 4.2548
Ge, ppm | 0.12194 | 0.024121 | 0.073702 | 0.17019 | 0.049581 | 0.19431 | 19.78% | 39.56% | 59.34% | 0.11585 | 0.12804
Hf, ppm | 0.34175 | 0.043431 | 0.25489 | 0.42861 | 0.21146 | 0.47204 | 12.71% | 25.42% | 38.13% | 0.32466 | 0.35884
Ho, ppm | 0.40563 | 0.032026 | 0.34157 | 0.46968 | 0.30955 | 0.50171 | 7.895% | 15.79% | 23.69% | 0.38535 | 0.42591
In, ppm | 0.066433 | 0.0046604 | 0.057112 | 0.075754 | 0.052452 | 0.080414 | 7.015% | 14.03% | 21.05% | 0.063112 | 0.069755
K, wt.% | 1.0152 | 0.062364 | 0.89044 | 1.1399 | 0.82808 | 1.2023 | 6.143% | 12.29% | 18.43% | 0.96441 | 1.0659
La, ppm | 19.2174 | 3.5252 | 12.167 | 26.2678 | 8.6418 | 29.793 | 18.34% | 36.69% | 55.03% | 18.2565 | 20.1782
Li, ppm | 44.4363 | 1.8253 | 40.7858 | 48.0868 | 38.9605 | 49.9121 | 4.108% | 8.215% | 12.32% | 42.2145 | 46.6581
Lu, ppm | 0.095156 | 0.0066971 | 0.081762 | 0.10855 | 0.075065 | 0.11525 | 7.038% | 14.08% | 21.11% | 0.090399 | 0.099914
Mg, wt.% | 0.80997 | 0.032079 | 0.74582 | 0.87413 | 0.71374 | 0.90621 | 3.961% | 7.921% | 11.88% | 0.76948 | 0.85047
Mn, wt.% | 0.032725 | 0.0011306 | 0.030463 | 0.034986 | 0.029333 | 0.036116 | 3.455% | 6.91% | 10.36% | 0.031088 | 0.034361
Mo, ppm | 2.3721 | 0.22048 | 1.9312 | 2.8131 | 1.7107 | 3.0336 | 9.295% | 18.59% | 27.88% | 2.2535 | 2.4907
Na, wt.% | 0.1593 | 0.024182 | 0.11094 | 0.20766 | 0.086754 | 0.23185 | 15.18% | 30.36% | 45.54% | 0.15134 | 0.16727
Nb, ppm | 0.78237 | 0.25849 | 0.26538 | 1.2994 | 0.0068906 | 1.5578 | 33.04% | 66.08% | 99.12% | 0.74325 | 0.82149
Ni, ppm | 18.6571 | 0.78725 | 17.0826 | 20.2316 | 16.2953 | 21.0188 | 4.22% | 8.439% | 12.66% | 17.7242 | 19.5899
P, wt.% | 0.064041 | 0.0027979 | 0.058445 | 0.069636 | 0.055647 | 0.072434 | 4.369% | 8.738% | 13.11% | 0.060839 | 0.067243
Pb, ppm | 4.794 | 0.50153 | 3.791 | 5.7971 | 3.2894 | 6.2986 | 10.46% | 20.92% | 31.38% | 4.5543 | 5.0337
Rb, ppm | 105.167 | 6.1252 | 92.9165 | 117.4175 | 86.7913 | 123.5428 | 5.824% | 11.65% | 17.47% | 99.9087 | 110.4254
S, wt.% | 0.255 | 0.016936 | 0.22112 | 0.28887 | 0.20419 | 0.3058 | 6.642% | 13.28% | 19.93% | 0.24225 | 0.26774
Sb, ppm | 0.24621 | 0.037106 | 0.172 | 0.32042 | 0.13489 | 0.35753 | 15.07% | 30.14% | 45.21% | 0.2339 | 0.25852
Sc, ppm | 8.1571 | 0.23305 | 7.691 | 8.6232 | 7.4579 | 8.8562 | 2.857% | 5.714% | 8.571% | 7.7492 | 8.5649
Se, ppm | 0.39741 | 0.12125 | 0.1549 | 0.63992 | 0.033645 | 0.76117 | 30.51% | 61.02% | 91.53% | 0.37754 | 0.41728
Sn, ppm | 3.5449 | 0.19935 | 3.1462 | 3.9436 | 2.9468 | 4.1429 | 5.624% | 11.25% | 16.87% | 3.3676 | 3.7221
Sr, ppm | 19.313 | 2.4339 | 14.4451 | 24.1809 | 12.0112 | 26.6148 | 12.6% | 25.21% | 37.81% | 18.3474 | 20.2787
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.57463 | 0.051161 | 0.4723 | 0.67695 | 0.42114 | 0.72811 | 8.903% | 17.81% | 26.71% | 0.5459 | 0.60336
Te, ppm | 0.049128 | 0.014241 | 0.020646 | 0.077611 | 0.0064044 | 0.091852 | 28.99% | 57.98% | 86.96% | 0.046672 | 0.051585
Th, ppm | 8.5151 | 1.6267 | 5.2617 | 11.7684 | 3.635 | 13.3951 | 19.1% | 38.21% | 57.31% | 8.0893 | 8.9408
Ti, wt.% | 0.27265 | 0.022752 | 0.22715 | 0.31816 | 0.2044 | 0.34091 | 8.345% | 16.69% | 25.03% | 0.25902 | 0.28629
Tl, ppm | 0.67857 | 0.028811 | 0.62094 | 0.73619 | 0.59213 | 0.765 | 4.246% | 8.492% | 12.74% | 0.64464 | 0.71249
U, ppm | 3.4653 | 0.28277 | 2.8997 | 4.0308 | 2.6169 | 4.3136 | 8.16% | 16.32% | 24.48% | 3.292 | 3.6385
V, ppm | 60.7629 | 2.44 | 55.883 | 65.6428 | 53.4431 | 68.0828 | 4.016% | 8.031% | 12.05% | 57.7248 | 63.8011
W, ppm | 10.1375 | 1.3551 | 7.4273 | 12.8476 | 6.0723 | 14.2027 | 13.37% | 26.73% | 40.1% | 9.6306 | 10.6443
Y, ppm | 10.8395 | 0.58342 | 9.6726 | 12.0063 | 9.0892 | 12.5897 | 5.382% | 10.76% | 16.15% | 10.2975 | 11.3815
Yb, ppm | 0.7082 | 0.017464 | 0.67327 | 0.74313 | 0.65581 | 0.76059 | 2.466% | 4.932% | 7.398% | 0.67279 | 0.74361
Zn, ppm | 77.0612 | 4.1612 | 68.7388 | 85.3836 | 64.5776 | 89.5448 | 5.4% | 10.8% | 16.2% | 73.2081 | 80.9143
Zr, ppm | 8.274 | 0.68123 | 6.9115 | 9.6364 | 6.2303 | 10.3177 | 8.233% | 16.47% | 24.7% | 7.8603 | 8.6877
Infrared Combustion
S, wt.% | 0.25716 | 0.014894 | 0.22737 | 0.28695 | 0.21248 | 0.30185 | 5.792% | 11.58% | 17.38% | 0.2443 | 0.27002
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 284
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 284
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Mineral Solutions (BVMS), Al Wadi District, Jeddah, Saudi Arabia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Minerals Ltd, Bibiani, Western North Region, Ghana
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Ghana Ltd, Obuasi, Ashanti, Ghana
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 284 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 11 | 6 | 261044 | 0.085054 | 0.41726 | 0.41656
2 | 1 | ANSLu | 1 | 2 | 1 | 261045 | 0.082628 | 0.38451 | 0.41482
2 | 1 | ANSLu | 1 | 3 | 2 | 261046 | 0.085451 | 0.40467 | 0.41589
2 | 1 | ANSLu | 1 | 15 | 8 | 261047 | 0.086103 | 0.39689 | 0.41548
2 | 1 | ANSLu | 1 | 8 | 4 | 261048 | 0.083629 | 0.4335 | 0.41743
2 | 1 | ANSLu | 1 | 18 | 9 | 261049 | 0.084023 | 0.40797 | 0.41607
2 | 1 | ANSLu | 1 | 13 | 7 | 261050 | 0.082955 | 0.43928 | 0.41773
2 | 1 | ANSLu | 1 | 20 | 10 | 261051 | 0.085336 | 0.43422 | 0.41746
2 | 1 | ANSLu | 1 | 12 | 6 | 261052 | 0.08829 | 0.42358 | 0.4169
2 | 1 | ANSLu | 1 | 6 | 3 | 261053 | 0.082804 | 0.42657 | 0.41706
2 | 1 | ANSLu | 1 | 7 | 4 | 261054 | 0.084282 | 0.43067 | 0.41727
2 | 1 | ANSLu | 1 | 14 | 7 | 261055 | 0.084688 | 0.40135 | 0.41571
2 | 1 | ANSLu | 1 | 1 | 1 | 261056 | 0.08602 | 0.40756 | 0.41604
2 | 1 | ANSLu | 1 | 4 | 2 | 261057 | 0.087214 | 0.43176 | 0.41733
2 | 1 | ANSLu | 1 | 5 | 3 | 261058 | 0.084421 | 0.40164 | 0.41573
2 | 1 | ANSLu | 1 | 16 | 8 | 261059 | 0.086634 | 0.41411 | 0.41639
2 | 1 | ANSLu | 1 | 10 | 5 | 261060 | 0.086866 | 0.43017 | 0.41725
2 | 1 | ANSLu | 1 | 17 | 9 | 261061 | 0.084552 | 0.41011 | 0.41618
2 | 1 | ANSLu | 1 | 9 | 5 | 261062 | 0.085562 | 0.41889 | 0.41665
2 | 1 | ANSLu | 1 | 19 | 10 | 261063 | 0.083874 | 0.41572 | 0.41648
Mean | [6 empty] | 0.085019 | 0.41652 | 0.41652
Median | [7 empty] | 0.41649 | 0.41652
SD | [7 empty] | 0.014687 | 0.00078185
RSD | [7 empty] | 3.526% | 0.1877%
RSD Upscale Factor | [7 empty] | 0.053235
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 284 (Certified Value 0.404 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 50g | 30g | 30g | 25g | 30g | 50g | 50g | 30g | 50g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 50g | 25g | 15g | 30g | 40g | 30g | 30g | 3
| 1 | 1 | 0.40756 | 0.42 | 0.3908 | 0.41825 | 0.405 | 0.40721 | 0.461 | 0.4 | 0.42 | 0.41 | 0.43 | 0.448 | 0.388 | 0.409 | 0.399 | 0.404 | 0.403 | 0.406 | 0.408 | 0.399 | 0.399 | 0.412 | 0.38 | 0.427 | 0.382 | 0.409 | 0.386 | 0.387 | 0.39 | 0.44 | 0.388 | 0.41 | 1
| 1 | 2 | 0.38451 | 0.38 | 0.3816 | 0.41559 | 0.401 | 0.37692 | 0.43 | 0.4 | 0.413 | 0.45 | 0.42 | 0.431 | 0.392 | 0.406 | 0.383 | 0.4 | 0.404 | 0.413 | 0.4117 | 0.408 | 0.406 | 0.413 | 0.38 | 0.428 | 0.376 | 0.399 | 0.383 | 0.4037 | 0.369 | 0.43 | 0.408 | 0.42 | 0
| 1 | 3 | 0.40467 | 0.4 | 0.3908 | 0.41592 | 0.391 | 0.42263 | 0.438 | 0.4 | 0.417 | 0.4 | 0.43 | 0.428 | 0.401 | 0.406 | 0.382 | 0.397 | 0.408 | 0.413 | 0.4199 | 0.406 | 0.404 | 0.413 | 0.39 | 0.428 | 0.367 | 0.408 | 0.396 | 0.386 | 0.393 | 0.44 | 0.405 | 0.4 | 16
| 1 | 4 | 0.43176 | 0.38 | 0.3984 | 0.41973 | 0.383 | 0.42943 | 0.436 | 0.41 | 0.4 | 0.41 | 0.42 | 0.41 | 0.415 | 0.405 | 0.391 | 0.398 | 0.406 | 0.419 | 0.41071 | 0.401 | 0.386 | 0.41 | 0.38 | 0.429 | 0.365 | 0.403 | 0.397 | 0.408 | 0.38 | 0.43 | 0.405 | 0.4 | 0.40389
| 1 | 5 | 0.40164 | 0.41 | 0.3816 | 0.41834 | 0.417 | 0.40748 | 0.44 | 0.4 | 0.398 | 0.44 | 0.43 | 0.403 | 0.41 | 0.41 | 0.391 | 0.401 | 0.4 | 0.42 | 0.41679 | 0.398 | 0.401 | 0.402 | 0.38 | 0.422 | 0.384 | 0.41 | 0.401 | 0.405 | 0.396 | 0.44 | 0.393 | 0.43 | 7
| 1 | 6 | 0.42657 | 0.4 | 0.4 | 0.41941 | 0.402 | 0.4068 | 0.456 | 0.4 | 0.413 | 0.45 | 0.42 | 0.404 | 0.41 | 0.396 | 0.4 | 0.402 | 0.405 | 0.419 | 0.41814 | 0.391 | 0.398 | 0.415 | 0.37 | 0.428 | 0.383 | 0.405 | 0.39 | 0.403 | 0.34 | 0.45 | 0.381 | 0.39
|  | 7 | 2.00: 0.43067
|  | 8 | 2.00: 0.4335
|  | 9 | 2.00: 0.41889
|  | 10 | 2.00: 0.43017
|  | 11 | 2.00: 0.41726
|  | 12 | 2.00: 0.42358
|  | 13 | 2.00: 0.43928
|  | 14 | 2.00: 0.40135
|  | 15 | 2.00: 0.39689
|  | 16 | 2.00: 0.41411
|  | 17 | 2.00: 0.41011
|  | 18 | 2.00: 0.40797
|  | 19 | 2.00: 0.41572
|  | 20 | 2.00: 0.43422
| Mean |  | 0.41652 | 0.39833 | 0.39053 | 0.41787 | 0.39983 | 0.40841 | 0.4435 | 0.40167 | 0.41017 | 0.42667 | 0.425 | 0.42067 | 0.40267 | 0.40533 | 0.391 | 0.40033 | 0.40433 | 0.415 | 0.41421 | 0.4005 | 0.399 | 0.41083 | 0.38 | 0.427 | 0.37617 | 0.40567 | 0.39217 | 0.39878 | 0.378 | 0.43833 | 0.39667 | 0.40833
| Median |  | 0.41649 | 0.4 | 0.3908 | 0.41829 | 0.4015 | 0.40735 | 0.439 | 0.4 | 0.413 | 0.425 | 0.425 | 0.419 | 0.4055 | 0.406 | 0.391 | 0.4005 | 0.4045 | 0.416 | 0.41425 | 0.4 | 0.4 | 0.4125 | 0.38 | 0.428 | 0.379 | 0.4065 | 0.393 | 0.40335 | 0.385 | 0.44 | 0.399 | 0.405
| Std Dev. |  | 0.014687 | 0.016021 | 0.0078899 | 0.0017433 | 0.011737 | 0.018114 | 0.012194 | 0.0040825 | 0.0090646 | 0.022509 | 0.0054772 | 0.017952 | 0.010875 | 0.0049666 | 0.0076158 | 0.002582 | 0.0027325 | 0.0054037 | 0.0047228 | 0.006091 | 0.0070427 | 0.0046224 | 0.0063246 | 0.0025298 | 0.0083766 | 0.0041793 | 0.006969 | 0.0096727 | 0.021081 | 0.0075277 | 0.010967 | 0.01472
| Rel.Std.Dev. |  | 3.526% | 4.022% | 2.02% | 0.4172% | 2.936% | 4.435% | 2.75% | 1.016% | 2.21% | 5.276% | 1.289% | 4.267% | 2.701% | 1.225% | 1.948% | 0.645% | 0.6758% | 1.302% | 1.14% | 1.521% | 1.765% | 1.125% | 1.664% | 0.5925% | 2.227% | 1.03% | 1.777% | 2.426% | 5.577% | 1.717% | 2.765% | 3.605%
| PDM3 |  | 3.128% | -1.375% | -3.306% | 3.463% | -1.004% | 1.12% | 9.808% | -0.5499% | 1.555% | 5.64% | 5.227% | 4.154% | -0.3023% | 0.3579% | -3.191% | -0.88% | 0.1103% | 2.751% | 2.555% | -0.8388% | -1.21% | 1.72% | -5.914% | 5.722% | -6.864% | 0.4405% | -2.902% | -1.264% | -6.41% | 8.529% | -1.788% | 1.101%
| Z-Score (Absolute) |  | NA | 0.53 | 1.21 | 1.19 | 0.4 | 0.36 | 3.44 | 0.23 | 0.51 | 1.96 | 1.82 | 1.44 | 0.15 | 0.09 | 1.17 | 0.35 | 0 | 0.94 | 0.87 | 0.34 | 0.47 | 0.57 | 2.14 | 1.99 | 2.48 | 0.12 | 1.07 | 0.49 | 2.32 | 2.99 | 0.67 | 0.35
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 284 (Certified Value 0.411 ppm) | [17 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | 5.05 | 5.06 | 5.07 | 5.08 | 5.09 | 5.11 | 5.12 | 5.13 | 5.14 | 5.15 | 5.16 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 330±15 g | 3
| 1 | 1 | 0.40756 | 0.377 | 0.438 | 0.404 | 0.399 | 0.4 | 0.405 | 0.373 | 0.44 | 0.42 | 0.39 | 0.37 | 0.42 | 0.43 | 0.41 | 0.42 | 1
| 1 | 2 | 0.38451 | 0.385 | 0.384 | 0.435 | 0.405 | 0.41 | 0.38 | 0.432 | 0.385 | 0.39 | 0.39 | 0.4 | 0.42 | 0.39 | 0.413 | 0.442 | 0
| 1 | 3 | 0.40467 | 0.337 | 0.43 | 0.416 | 0.419 | 0.412 | 0.418 | 0.408 | 0.467 | 0.56 | 0.4 | 0.37 | 0.38 | 0.43 | 0.417 | 0.416 | 16
| 1 | 4 | 0.43176 | 0.355 | 0.428 | 0.424 | 0.414 | 0.406 | 0.4 | 0.373 | 0.373 | 0.56 | 0.47 | 0.43 | 0.39 | 0.4 | 0.421 | 0.434 | 0.4109
| 1 | 5 | 0.40164 | 0.413 | 0.402 | 0.644 | 0.411 | 0.444 | 0.416 | 0.403 | 0.397 | 0.38 | 0.45 | 0.44 | 0.46 | 0.41 | 0.664 | 0.356 | 9
| 1 | 6 | 0.42657 | 0.443 | 0.412 | 0.722 | 0.4 | 0.428 | 0.351 | 0.374 | 0.427 | 0.35 | 0.38 | 0.39 | 0.47 | 0.38 | 0.651 | 0.409
|  | 7 | 2.00: 0.43067
|  | 8 | 2.00: 0.4335
|  | 9 | 2.00: 0.41889
|  | 10 | 2.00: 0.43017
|  | 11 | 2.00: 0.41726
|  | 12 | 2.00: 0.42358
|  | 13 | 2.00: 0.43928
|  | 14 | 2.00: 0.40135
|  | 15 | 2.00: 0.39689
|  | 16 | 2.00: 0.41411
|  | 17 | 2.00: 0.41011
|  | 18 | 2.00: 0.40797
|  | 19 | 2.00: 0.41572
|  | 20 | 2.00: 0.43422
| Mean |  | 0.41652 | 0.385 | 0.41567 | 0.5075 | 0.408 | 0.41667 | 0.395 | 0.39383 | 0.41483 | 0.44333 | 0.41333 | 0.4 | 0.42333 | 0.40667 | 0.496 | 0.41283
| Median |  | 0.41649 | 0.381 | 0.42 | 0.4295 | 0.408 | 0.411 | 0.4025 | 0.3885 | 0.412 | 0.405 | 0.395 | 0.395 | 0.42 | 0.405 | 0.419 | 0.418
| Std Dev. |  | 0.014687 | 0.038512 | 0.020294 | 0.13853 | 0.008 | 0.01633 | 0.025519 | 0.024506 | 0.035947 | 0.093095 | 0.037238 | 0.029665 | 0.036148 | 0.020656 | 0.12522 | 0.030347
| Rel.Std.Dev. |  | 3.526% | 10% | 4.882% | 27.3% | 1.961% | 3.919% | 6.46% | 6.223% | 8.665% | 21% | 9.009% | 7.416% | 8.539% | 5.079% | 25.25% | 7.351%
| PDM3 |  | 1.368% | -6.303% | 1.16% | 23.51% | -0.7058% | 1.403% | -3.87% | -4.153% | 0.9572% | 7.893% | 0.5922% | -2.653% | 3.026% | -1.03% | 20.71% | 0.4705%
| Z-Score (Absolute) |  | NA | 1.91 | 0.16 | 6.35 | 0.36 | 0.22 | 1.24 | 1.31 | 0.1 | 2.02 | 0 | 0.9 | 0.67 | 0.45 | 5.57 | 0.03
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 284 (Certified Value 0.394 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.04 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.28 | 1.32 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | 3
|  |  | 0.085g | 50g | 25g | 50g | 30g | 50g | 30g | 10g | 15g | 25g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 25g | 15g | 25g | 10g | 50g | 30g | 3
| 1 | 1 | 0.40756 | 0.39 | 0.402 | 0.4 | 0.368 | 0.48 | 0.4 | 0.4 | 0.384 | 0.41 | 0.375 | 0.398 | 0.427 | 0.383 | 0.39083 | 0.357 | 0.4 | 0.34 | 0.3821 | 0.405 | 0.361 | 0.408 | 0.38 | 1
| 1 | 2 | 0.38451 | 0.32 | 0.404 | 0.4 | 0.373 | 0.4 | 0.4 | 0.39 | 0.3729 | 0.424 | 0.369 | 0.408 | 0.419 | 0.419 | 0.41437 | 0.365 | 0.4 | 0.31 | 0.3856 | 0.422 | 0.374 | 0.39 | 0.41 | 0
| 1 | 3 | 0.40467 | 0.38 | 0.405 | 0.4 | 0.377 | 0.48 | 0.42 | 0.4 | 0.3725 | 0.42 | 0.368 | 0.405 | 0.42 | 0.388 | 0.39921 | 0.356 | 0.39 | 0.31 | 0.3859 | 0.419 | 0.372 | 0.4 | 0.38 | 16
| 1 | 4 | 0.43176 | 0.34 | 0.418 | 0.38 | 0.375 | 0.4 | 0.41 | 0.38 | 0.3691 | 0.407 | 0.372 | 0.405 | 0.445 | 0.41 | 0.39645 | 0.347 | 0.38 | 0.32 | 0.414 | 0.406 | 0.374 | 0.418 | 0.38 | 0.39429
| 1 | 5 | 0.40164 | 0.39 | 0.405 | 0.4 | 0.378 | 0.46 | 0.4 | 0.4 | 0.3915 | 0.419 | 0.372 | 0.408 | 0.428 | 0.405 | 0.41555 | 0.378 | 0.4 | 0.31 | 0.4088 | 0.4 | 0.378 | 0.403 | 0.37 | 11
| 1 | 6 | 0.42657 | 0.39 | 0.402 | 0.4 | 0.372 | 0.48 | 0.41 | 0.4 | 0.3916 | 0.423 | 0.381 | 0.408 | 0.434 | 0.411 | 0.39157 | 0.375 | 0.39 | 0.31 | 0.3712 | 0.411 | 0.393 | 0.418 | 0.37
|  | 7 | 2.00: 0.43067
|  | 8 | 2.00: 0.4335
|  | 9 | 2.00: 0.41889
|  | 10 | 2.00: 0.43017
|  | 11 | 2.00: 0.41726
|  | 12 | 2.00: 0.42358
|  | 13 | 2.00: 0.43928
|  | 14 | 2.00: 0.40135
|  | 15 | 2.00: 0.39689
|  | 16 | 2.00: 0.41411
|  | 17 | 2.00: 0.41011
|  | 18 | 2.00: 0.40797
|  | 19 | 2.00: 0.41572
|  | 20 | 2.00: 0.43422
| Mean |  | 0.41652 | 0.36833 | 0.406 | 0.39667 | 0.37383 | 0.45 | 0.40667 | 0.395 | 0.38027 | 0.41717 | 0.37283 | 0.40533 | 0.42883 | 0.40267 | 0.40133 | 0.363 | 0.39333 | 0.31667 | 0.39127 | 0.4105 | 0.37533 | 0.40617 | 0.38167
| Median |  | 0.41649 | 0.385 | 0.4045 | 0.4 | 0.374 | 0.47 | 0.405 | 0.4 | 0.37845 | 0.4195 | 0.372 | 0.4065 | 0.4275 | 0.4075 | 0.39783 | 0.361 | 0.395 | 0.31 | 0.38575 | 0.4085 | 0.374 | 0.4055 | 0.38
| Std Dev. |  | 0.014687 | 0.030605 | 0.0060332 | 0.008165 | 0.003656 | 0.039497 | 0.008165 | 0.0083666 | 0.010078 | 0.0070261 | 0.0047081 | 0.0038816 | 0.0096626 | 0.014123 | 0.011011 | 0.01195 | 0.008165 | 0.012111 | 0.016564 | 0.0085499 | 0.010386 | 0.010889 | 0.01472
| Rel.Std.Dev. |  | 3.526% | 8.309% | 1.486% | 2.058% | 0.978% | 8.777% | 2.008% | 2.118% | 2.65% | 1.684% | 1.263% | 0.9576% | 2.253% | 3.507% | 2.744% | 3.292% | 2.076% | 3.824% | 4.233% | 2.083% | 2.767% | 2.681% | 3.857%
| PDM3 |  | 5.637% | -6.584% | 2.969% | 0.6019% | -5.189% | 14.13% | 3.138% | 0.1792% | -3.557% | 5.801% | -5.443% | 2.8% | 8.76% | 2.124% | 1.785% | -7.937% | -0.2435% | -19.69% | -0.7676% | 4.11% | -4.809% | 3.011% | -3.202%
| Z-Score (Absolute) |  | NA | 1.29 | 0.48 | 0.04 | 1.03 | 2.53 | 0.51 | 0.04 | 0.73 | 1 | 1.08 | 0.44 | 1.54 | 0.32 | 0.26 | 1.54 | 0.12 | 3.7 | 0.21 | 0.69 | 0.96 | 0.48 | 0.66
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 284 (Certified Value 0.385 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.31 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 20g | 15g | 05g | 05g | 15g | 200g | 3
| 1 | 1 | 0.40756 | 0.36 | 0.35 | 0.41161 | 0.4 | 0.41 | 0.4 | 0.44353 | 0.39 | 0.445 | 0.42 | 0.39 | 0.42 | 0.42 | 0.35 | 0.38 | 0.304 | 0.32 | 0.38 | 0.4 | 0.38 | 0.38 | 1
| 1 | 2 | 0.38451 | 0.31 | 0.36 | 0.40489 | 0.4 | 0.37 | 0.42 | 0.42628 | 0.39 | 0.431 | 0.42 | 0.39 | 0.41 | 0.44 | 0.3 | 0.35 | 0.38516 | 0.36 | 0.4 | 0.4 | 0.38 | 0.326 | 0
| 1 | 3 | 0.40467 | 0.37 | 0.35 | 0.41004 | 0.4 | 0.35821 | 0.4 | 0.4397 | 0.41 | 0.434 | 0.42 | 0.39 | 0.41 | 0.44 | 0.36 | 0.3 | 0.37538 | 0.34 | 0.4 | 0.4 | 0.35 | 0.332 | 16
| 1 | 4 | 0.43176 | 0.32 | 0.36 | 0.41266 | 0.4 | 0.36453 | 0.41 | 0.43969 | 0.43 | 0.425 | 0.43 | 0.39 | 0.42 | 0.43 | 0.33 | 0.33 | 0.38323 | 0.34 | 0.38 | 0.4 | 0.38 | 0.238 | 0.38534
| 1 | 5 | 0.40164 | 0.37 | 0.35 | 0.40835 | 0.4 | 0.38614 | 0.43 | 0.42822 | 0.39 | 0.445 | 0.43 | 0.39 | 0.41 | 0.42 | 0.35 | 0.36 | 0.37919 | 0.34 | 0.4 | 0.4 | 0.35 | 0.251 | 13
| 1 | 6 | 0.42657 | 0.37 | 0.35 | 0.39887 | 0.39 | 0.36816 | 0.41 | 0.43227 | 0.39 | 0.427 | 0.43 | 0.4 | 0.4 | 0.42 | 0.33 | 0.31 | 0.373 | 0.33 | 0.38 | 0.4 | 0.37 | 0.292
|  | 7 | 2.00: 0.43067
|  | 8 | 2.00: 0.4335
|  | 9 | 2.00: 0.41889
|  | 10 | 2.00: 0.43017
|  | 11 | 2.00: 0.41726
|  | 12 | 2.00: 0.42358
|  | 13 | 2.00: 0.43928
|  | 14 | 2.00: 0.40135
|  | 15 | 2.00: 0.39689
|  | 16 | 2.00: 0.41411
|  | 17 | 2.00: 0.41011
|  | 18 | 2.00: 0.40797
|  | 19 | 2.00: 0.41572
|  | 20 | 2.00: 0.43422
| Mean |  | 0.41652 | 0.35 | 0.35333 | 0.40774 | 0.39833 | 0.37617 | 0.41167 | 0.43495 | 0.4 | 0.4345 | 0.425 | 0.39167 | 0.41167 | 0.42833 | 0.33667 | 0.33833 | 0.36666 | 0.33833 | 0.39 | 0.4 | 0.36833 | 0.30317
| Median |  | 0.41649 | 0.365 | 0.35 | 0.40919 | 0.4 | 0.36908 | 0.41 | 0.43598 | 0.39 | 0.4325 | 0.425 | 0.39 | 0.41 | 0.425 | 0.34 | 0.34 | 0.37728 | 0.34 | 0.39 | 0.4 | 0.375 | 0.309
| Std Dev. |  | 0.014687 | 0.027568 | 0.005164 | 0.0051283 | 0.0040825 | 0.018995 | 0.01169 | 0.007019 | 0.016733 | 0.0087121 | 0.0054772 | 0.0040825 | 0.0075277 | 0.0098319 | 0.021602 | 0.030605 | 0.031036 | 0.013292 | 0.010954 | 6.0809e-17 | 0.01472 | 0.05357
| Rel.Std.Dev. |  | 3.526% | 7.877% | 1.462% | 1.258% | 1.025% | 5.05% | 2.84% | 1.614% | 4.183% | 2.005% | 1.289% | 1.042% | 1.829% | 2.295% | 6.417% | 9.046% | 8.465% | 3.929% | 2.809% | 1.52e-14% | 3.996% | 17.67%
| PDM3 |  | 8.093% | -9.17% | -8.305% | 5.813% | 3.373% | -2.378% | 6.833% | 12.88% | 3.806% | 12.76% | 10.29% | 1.643% | 6.833% | 11.16% | -12.63% | -12.2% | -4.847% | -12.2% | 1.21% | 3.806% | -4.412% | -21.32%
| Z-Score (Absolute) |  | NA | 0.94 | 0.87 | 0.39 | 0.17 | 0.34 | 0.48 | 1.02 | 0.21 | 1.01 | 0.79 | 0.02 | 0.48 | 0.87 | 1.25 | 1.21 | 0.56 | 1.21 | 0.02 | NA | 0.52 | 2.03
Results from laboratory 1.22 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 284 (Certified Value 0.095 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.15336 | 0.1 | 0.4 | < 0.1 | < 0.5 | < 0.3 | 0.03 | 0.08 | 0.09 | 0.09 | 0.09 | 0.06 | < 0.1 | < 0.1 | 0.21 | 0.1 | 0.09 | < 0.1 | 0.1 | 0.1 | 0.11 | 2 | 0.26 | 0.1 |  |  | 1
| 1 | 2 | 0.15341 | 0.09 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.09 | 0.1 | 0.09 | 0.09 | 0.11 | 0.09 | < 0.1 | < 0.1 | 0.22 | 0.1 | 0.08 | < 0.1 | 0.11 | 0.09 | 0.09 | 2 | 0.24 | 0.1 |  |  | 22
| 1 | 3 | 0.14991 | 0.1 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.04 | 0.08 | 0.09 | 0.1 | 0.09 | 0.04 | 0.10001 | < 0.1 | 0.19 | 0.1 | 0.09 | < 0.1 | 0.1 | 0.08 | 0.13 | 2.1 | 0.219 | 0.1 |  |  | 16
| 1 | 4 | 0.1688 | 0.1 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.05 | 0.06 | 0.12 | 0.09 | 0.1 | 0.04 | 0.1012 | < 0.1 | 0.19 | < 0.1 | 0.09 | < 0.1 | 0.1 | 0.09 | 0.09 | 2.4 | 0.24 | 0.1 |  |  | 0.095207
| 1 | 5 | 0.1435 | 0.1 | 0.4 | < 0.1 | < 0.5 | < 0.3 | 0.03 | 0.1 | 0.13 | 0.09 | 0.07 | 0.05 | < 0.1 | < 0.1 | 0.22 | 0.1 | 0.08 | 0.1 | 0.1 | 0.1 | 0.08 | 2 | 0.22 | 0.1 |  |  | 15
| 1 | 6 | 0.13768 | 0.11 | 0.4 | < 0.1 | < 0.5 | < 0.3 | 0.07 | 0.07 | 0.1 | 0.09 | 0.08 | 0.02 | < 0.1 | < 0.1 | 0.2 | 0.1 | 0.09 | < 0.1 | 0.11 | 0.09 | 0.1 | 2.2 | 0.22 | 0.09
| Mean |  | 0.15111 | 0.1 | 0.4 |  |  |  | 0.051667 | 0.081667 | 0.10333 | 0.091667 | 0.09 | 0.05 | 0.10061 |  | 0.205 | 0.1 | 0.086667 | 0.1 | 0.10333 | 0.091667 | 0.1 | 2.1167 | 0.23317 | 0.098333
| Median |  | 0.15164 | 0.1 | 0.4 |  |  |  | 0.045 | 0.08 | 0.095 | 0.09 | 0.09 | 0.045 | 0.10061 |  | 0.205 | 0.1 | 0.09 | 0.1 | 0.1 | 0.09 | 0.095 | 2.05 | 0.23 | 0.1
| Std Dev. |  | 0.01061 | 0.0063246 | 6.7987e-17 |  |  |  | 0.024014 | 0.016021 | 0.017512 | 0.0040825 | 0.014142 | 0.023664 | 0.00084081 |  | 0.013784 | 0 | 0.005164 |  | 0.005164 | 0.0075277 | 0.017889 | 0.16021 | 0.016497 | 0.0040825
| Rel.Std.Dev. |  | 7.021% | 6.325% | 1.7e-14% |  |  |  | 46.48% | 19.62% | 16.95% | 4.454% | 15.71% | 47.33% | 0.8358% |  | 6.724% | 0% | 5.958% |  | 4.997% | 8.212% | 17.89% | 7.569% | 7.075% | 4.152%
| PDM3 |  | 58.72% | 5.035% | 320.1% |  |  |  | -45.73% | -14.22% | 8.536% | -3.718% | -5.469% | -47.48% | 5.671% |  | 115.3% | 5.035% | -8.97% | 5.035% | 8.536% | -3.718% | 5.035% | 2123% | 144.9% | 3.284%
| Z-Score (Absolute) |  | 1.82 | 0.16 | NA | 1.46 | 5.04 | 1.79 | 1.41 | 0.43 | 0.27 | 0.11 | 0.16 | 1.46 | 0.91 | 1.46 | 3.58 | NA | 0.27 | NA | 0.27 | 0.11 | 0.16 | NA | 4.49 | 0.11
Results from laboratories 1.09, 1.22, 1.24 and 1.29 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 284 (Certified Value 7.74 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 7.7964 | 7.5274 | 7.8455 | 7.49 | 7.6893 | 7.91 | 7.2876 | 7.66 | 7.46 | 8.28 | 7.53 | 6.97 | 7.34 | 7.9901 | 6.49 | 6.98 | 7.32 | 7.75 | 7.42 | 7.56 | 7.9907 | 7.7828 | 8.31 | 7.63 | 5.66 |  | 1
| 1 | 2 | 7.8293 | 7.6212 | 7.8506 | 7.61 | 7.5723 | 7.93 | 7.3625 | 7.75 | 7.7 | 8.06 | 7.49 | 7.4 | 7.86 | 8.0224 | 6.41 | 7.22 | 7.48 | 7.6 | 7.6 | 8.01 | 8.0048 | 7.9798 | 8.41 | 7.81 | 5.4 |  | 0
| 1 | 3 | 7.7968 | 7.7204 | 8.0945 | 7.57 | 7.5828 | 7.87 | 7.2505 | 8.06 | 7.57 | 7.96 | 7.4 | 7.35 | 7.68 | 8.0119 | 6.21 | 7.37 | 7.59 | 7.54 | 7.51 | 7.66 | 8.0167 | 7.9283 | 8.2 | 7.69 | 5.34 |  | 16
| 1 | 4 | 7.7398 | 7.7958 | 7.9489 | 7.48 | 7.722 | 7.92 | 7.4971 | 7.65 | 7.51 | 7.99 | 7.83 | 7.52 | 7.67 | 7.9734 | 6.18 | 7.21 | 7.38 | 7.75 | 7.58 | 7.99 | 8.1309 | 7.9679 | 8.26 | 8.07 | 5.76 |  | 7.7416
| 1 | 5 | 7.8053 | 7.5898 | 7.9588 | 7.59 | 7.4419 | 7.91 | 7.1851 | 7.96 | 7.78 | 8.13 | 7.61 | 7.34 | 7.74 | 7.9824 | 6.18 | 7.37 | 7.59 | 7.6 | 7.81 | 8.02 | 8.2404 | 8.0788 | 8.47 | 7.96 | 5.28 |  | 16
| 1 | 6 | 7.7805 | 7.6802 | 8.1112 | 7.46 | 7.5777 | 7.92 | 7.3052 | 7.41 | 7.59 | 8.13 | 7.56 | 7.35 | 7.84 | 7.9619 | 6.1 | 7.27 | 7.66 | 7.57 | 7.67 | 8 | 8.1348 | 8.033 | 8.68 | 8.33 | 5.08
| Mean |  | 7.7914 | 7.6558 | 7.9683 | 7.5333 | 7.5977 | 7.91 | 7.3147 | 7.7483 | 7.6017 | 8.0917 | 7.57 | 7.3217 | 7.6883 | 7.9904 | 6.2617 | 7.2367 | 7.5033 | 7.635 | 7.5983 | 7.8733 | 8.0864 | 7.9618 | 8.3883 | 7.915 | 5.42
| Median |  | 7.7966 | 7.6507 | 7.9539 | 7.53 | 7.5803 | 7.915 | 7.2964 | 7.705 | 7.58 | 8.095 | 7.545 | 7.35 | 7.71 | 7.9863 | 6.195 | 7.245 | 7.535 | 7.6 | 7.59 | 7.995 | 8.0738 | 7.9738 | 8.36 | 7.885 | 5.37
| Std Dev. |  | 0.029891 | 0.096362 | 0.11467 | 0.064083 | 0.09938 | 0.020976 | 0.10705 | 0.23404 | 0.11923 | 0.11583 | 0.14574 | 0.18498 | 0.18809 | 0.023019 | 0.1525 | 0.1439 | 0.13337 | 0.091815 | 0.13408 | 0.20666 | 0.098706 | 0.10226 | 0.17337 | 0.26121 | 0.25108
| Rel.Std.Dev. |  | 0.3836% | 1.259% | 1.439% | 0.8507% | 1.308% | 0.2652% | 1.463% | 3.021% | 1.569% | 1.431% | 1.925% | 2.526% | 2.446% | 0.2881% | 2.435% | 1.988% | 1.777% | 1.203% | 1.765% | 2.625% | 1.221% | 1.284% | 2.067% | 3.3% | 4.632%
| PDM3 |  | 0.6432% | -1.108% | 2.928% | -2.69% | -1.859% | 2.176% | -5.514% | 0.08734% | -1.807% | 4.522% | -2.216% | -5.424% | -0.6877% | 3.214% | -19.12% | -6.522% | -3.077% | -1.377% | -1.85% | 1.702% | 4.454% | 2.844% | 8.354% | 2.24% | -29.99%
| Z-Score (Absolute) |  | 0.36 | 0 | 0.83 | 0.32 | 0.15 | 0.67 | 0.9 | 0.25 | 0.14 | 1.16 | 0.23 | 0.89 | 0.09 | 0.89 | 3.7 | 1.11 | 0.4 | 0.06 | 0.15 | 0.58 | 1.14 | 0.81 | 1.94 | 0.69 | 5.93
| Analytical results for As in OREAS 284 (Certified Value 180 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 171.5923 | 190.1 | 162 | 231.6757 | 181 | 187 | 182.2 | 176 | 185 | 170.5 | 169 | 172 | 184.1348 | 26.7 | 174 | 178 | 186 | 178 | 173 | 181 | 173.3 | 185 | 202 | 171 | 176 |  | 1
| 1 | 2 | 173.6517 | 188.6 | 161 | 166.1341 | 185 | 190.2 | 176.2 | 181 | 182 | 159 | 176.5 | 182.5 | 182.038 | 30 | 173 | 179 | 186 | 157 | 187 | 181 | 176.7 |  | 187 | 177 | 179 |  | 12
| 1 | 3 | 175.7045 | 189.6 | 167 | 167.2504 | 187 | 195.3 | 186.7 | 175.5 | 179.5 | 160.5 | 183 | 180.5 | 179.8082 | 32.9 | 180 | 182 | 186 | 166 | 175 | 179 | 177.5 | 169.33 | 191 | 142 | 175 |  | 16
| 1 | 4 | 178.7258 | 189.8 | 159 | 210.1563 | 184 | 201.1 | 173.5 | 178.5 | 183 | 167 | 182.5 | 181.5 | 182.5845 | 27.9 | 181 | 181 | 190 | 164 | 183 | 174 | 179.4 | 175.95 | 185 | 180 | 178 |  | 179.9773
| 1 | 5 | 176.4759 | 191.1 | 179 | 174.1781 | 180 | 195.5 | 182.3 | 182.5 | 184 | 163.5 | 176 | 179 | 183.0608 | 28.5 | 177 | 180 | 190 | 189 | 192 | 185 | 177 | 188.85 | 177 | 161 | 180 |  | 17
| 1 | 6 | 177.2993 | 189.3 | 152 | 157.0641 | 186 | 199 | 176.9 | 181 | 187 | 157 | 181 | 182.5 | 180.4548 | 28.3 | 183 | 184 | 189 | 171 | 187 | 184 | 179.3 | 192.34 | 182 | 187 | 180
| Mean |  | 175.5749 | 189.75 | 163.3333 | 184.4098 | 183.8333 | 194.6833 | 179.6333 | 179.0833 | 183.4167 | 162.9167 | 178 | 179.6667 | 182.0135 | 29.05 | 178 | 180.6667 | 187.8333 | 170.8333 | 182.8333 | 180.6667 | 177.2 | 182.294 | 187.3333 | 169.6667 | 178
| Median |  | 176.0902 | 189.7 | 161.5 | 170.7142 | 184.5 | 195.4 | 179.55 | 179.75 | 183.5 | 162 | 178.75 | 181 | 182.3112 | 28.4 | 178.5 | 180.5 | 187.5 | 168.5 | 185 | 181 | 177.25 | 185 | 186 | 174 | 178.5
| Std Dev. |  | 2.5803 | 0.83606 | 9.0921 | 29.5901 | 2.7869 | 5.2913 | 4.9094 | 2.8882 | 2.5771 | 5.1129 | 5.3198 | 3.9833 | 1.6258 | 2.1649 | 4 | 2.1602 | 2.0412 | 11.3387 | 7.4409 | 3.9328 | 2.2289 | 9.4807 | 8.5946 | 16.1452 | 2.0976
| Rel.Std.Dev. |  | 1.47% | 0.4406% | 5.567% | 16.05% | 1.516% | 2.718% | 2.733% | 1.613% | 1.405% | 3.138% | 2.989% | 2.217% | 0.8932% | 7.452% | 2.247% | 1.196% | 1.087% | 6.637% | 4.07% | 2.177% | 1.258% | 5.201% | 4.588% | 9.516% | 1.178%
| PDM3 |  | -2.446% | 5.43% | -9.248% | 2.463% | 2.143% | 8.171% | -0.1911% | -0.4967% | 1.911% | -9.479% | -1.099% | -0.1726% | 1.131% | -83.86% | -1.099% | 0.3831% | 4.365% | -5.081% | 1.587% | 0.3831% | -1.543% | 1.287% | 4.087% | -5.729% | -1.099%
| Z-Score (Absolute) |  | 0.74 | 1.81 | 2.94 | 0.85 | 0.75 | 2.7 | 0.01 | 0.1 | 0.67 | 3.01 | 0.3 | 0 | 0.42 | 27.08 | 0.3 | 0.18 | 1.47 | 1.59 | 0.57 | 0.18 | 0.44 | 0.47 | 1.38 | 1.8 | 0.3
| Analytical results for B in OREAS 284 (Indicative Value 9.65 ppm) | [27 empty] | Indicative
B | Round | Replicate | 1.03: Lab; 1.04: Lab |  | 1
| No | No | 1.03: 1.10; 1.04: 1.11 |  | ppm
|  |  | 1.03: 4A*OES; 1.04: 4A*OES |  | 2
2
| 1 | 1 | 1.03: 2.1999; 1.04: 17 |  | 1
| 1 | 2 | 1.03: 2.9646; 1.04: 17 |  | 1
| 1 | 3 | 1.03: 3.2685; 1.04: 17 |  | 16
| 1 | 4 | 1.03: 3.2792; 1.04: 15 |  | 9.6451
| 1 | 5 | 1.03: 2.6594; 1.04: 16 |  | 7
| 1 | 6 | 1.03: 3.3702; 1.04: 16
| Mean |  | 1.03: 2.957; 1.04: 16.3333
| Median |  | 1.03: 3.1166; 1.04: 16.5
| Std Dev. |  | 1.03: 0.45462; 1.04: 0.8165
| Rel.Std.Dev. |  | 1.03: 15.37%; 1.04: 4.999%
| PDM3 |  | 1.03: -69.34%; 1.04: 69.34%
| Z-Score (Absolute) |  | 1.03: 0.67; 1.04: 0.67
| Analytical results for Ba in OREAS 284 (Certified Value 1045 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 1007.4985 | 1077.6 | 1100 | 996.348 | 980 | 1108 | 1087 | 1030 | 1090 | 1000 | 1010 | 1010 | 1097.8708 | > 1000 | 976 | 989 | 1080 | 1095 | 976 | 1055 | 1020.6 | 1090 | 950 | 1008 |  |  | 1
| 1 | 2 | 1037.5778 | 1075.6 | 1100 | 982.1612 | 987 | 1124 | 1033 | 1050 | 1070 | 1000 | 1050 | 1080 | 1113.8717 | > 1000 | 967 | 986 | 1050 | 1090 | 1031 | 1048 | 1050.7 | 1080 | 960 | 1007 |  |  | 13
| 1 | 3 | 1036.707 | 1064.8 | 1100 | 984.3104 | 973 | 1145 | 1126 | 1020 | 1140 | 990 | 1050 | 1060 | 1087.1807 | > 1000 | 984 | 966 | 1080 | 1074 | 990 | 1044 | 1058.8 | 1040 | 930 | 1022 |  |  | 16
| 1 | 4 | 1052.5851 | 1052.7 | 1100 | 1001.0564 | 980 | 1163 | 1066 | 1040 | 1110 | 1030 | 1080 | 1060 | 1097.8156 | > 1000 | 989 | 962 | 1080 | 1068 | 1029 | 1030 | 1042.6 | 1050 | 960 | 1056 |  |  | 1044.5144
| 1 | 5 | 1021.4504 | 1054.6 | 1110 | 970.9077 | 970 | 1152 | 1111 | 1050 | 1080 | 1020 | 1060 | 1070 | 1087.9828 | > 1000 | 997 | 980 | 1100 | 1075 | 1037 | 1075 | 1059.5 | 1080 | 950 | 1029 |  |  | 18
| 1 | 6 | 1027.7524 | 1062.3 | 1100 | 986.8055 | 986 | 1164 | 1062 | 1040 | 1080 | 990 | 1050 | 1090 | 1107.8201 | > 1000 | 1005 | 982 | 1070 | 1033 | 1032 | 1048 | 1058.3 | 1060 | 980 | 1034
| Mean |  | 1030.5952 | 1064.6 | 1101.6667 | 986.9315 | 979.3333 | 1142.6667 | 1080.8333 | 1038.3333 | 1095 | 1005 | 1050 | 1061.6667 | 1098.757 |  | 986.3333 | 977.5 | 1076.6667 | 1072.5 | 1015.8333 | 1050 | 1048.4167 | 1066.6667 | 955 | 1026
| Median |  | 1032.2297 | 1063.55 | 1100 | 985.558 | 980 | 1148.5 | 1076.5 | 1040 | 1085 | 1000 | 1050 | 1065 | 1097.8432 |  | 986.5 | 981 | 1080 | 1074.5 | 1030 | 1048 | 1054.5 | 1070 | 955 | 1025.5
| Std Dev. |  | 15.4585 | 10.3628 | 4.0825 | 10.72 | 6.802 | 22.4113 | 34.2077 | 11.6905 | 25.8844 | 16.4317 | 22.8035 | 27.8687 | 10.6053 |  | 13.8227 | 10.9864 | 16.3299 | 21.9157 | 25.9493 | 14.7919 | 15.1073 | 19.6638 | 16.4317 | 18.2975
| Rel.Std.Dev. |  | 1.5% | 0.9734% | 0.3706% | 1.086% | 0.6946% | 1.961% | 3.165% | 1.126% | 2.364% | 1.635% | 2.172% | 2.625% | 0.9652% |  | 1.401% | 1.124% | 1.517% | 2.043% | 2.554% | 1.409% | 1.441% | 1.843% | 1.721% | 1.783%
| PDM3 |  | -1.333% | 1.923% | 5.472% | -5.513% | -6.24% | 9.397% | 3.477% | -0.5918% | 4.833% | -3.783% | 0.5252% | 1.642% | 5.193% |  | -5.57% | -6.416% | 3.078% | 2.679% | -2.746% | 0.5252% | 0.3736% | 2.121% | -8.57% | -1.773%
| Z-Score (Absolute) |  | 0.42 | 0.32 | 1.13 | 1.38 | 1.55 | 2.03 | 0.67 | 0.26 | 0.98 | 0.98 | 0 | 0.26 | 1.07 | NA | 1.39 | 1.59 | 0.58 | 0.49 | 0.75 | 0 | 0.03 | 0.36 | 2.08 | 0.52
| Analytical results for Be in OREAS 284 (Certified Value 2.56 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.513 | 2.77 | 2.8 | 1.9583 | < 5 | 2.51 | 3 | 2.57 | 2.61 | 2.47 | 2.44 | 2.48 | 2.5153 | 2 | 2.5 | 2.3 | 2.54 | 3 | 2.64 | 2.47 | 2.68 | 3.1 | 3.07 | 2.3 |  |  | 1
| 1 | 2 | 2.526 | 2.88 | 2.6 | 1.8856 | < 5 | 2.6 | 3 | 2.62 | 2.57 | 2.46 | 2.55 | 2.62 | 2.4401 | 2 | 2.5 | 2.3 | 2.23 | 3 | 2.81 | 2.59 | 2.63 | 3.1 | 2.37 | 2.5 |  |  | 25
| 1 | 3 | 2.5626 | 2.73 | 2.8 | 1.9162 | < 5 | 2.59 | 2 | 2.56 | 2.56 | 2.42 | 2.57 | 2.59 | 2.3916 | 2 | 2.4 | 2.3 | 2.51 | 3 | 2.67 | 2.58 | 2.69 | 3 | 2.34 | 2.3 |  |  | 16
| 1 | 4 | 2.574 | 2.81 | 2.6 | 2.0204 | < 5 | 2.64 | 2 | 2.59 | 2.56 | 2.51 | 2.62 | 2.6 | 2.5928 | 2 | 2.5 | 2.2 | 2.34 | 3 | 2.74 | 2.53 | 2.76 | 3 | 3.05 | 2.2 |  |  | 2.56
| 1 | 5 | 2.5291 | 2.72 | 2.6 | 1.962 | < 5 | 2.16 | 2 | 2.64 | 2.62 | 2.47 | 2.56 | 2.59 | 2.5315 | 2 | 2.6 | 2.3 | 2.43 | 3 | 2.8 | 2.67 | 2.65 | 3.1 | 3.1 | 2.3 |  |  | 19
| 1 | 6 | 2.554 | 2.7 | 2.7 | 1.9889 | < 5 | 2.59 | 2 | 2.57 | 2.6 | 2.43 | 2.52 | 2.59 | 2.6214 | 2 | 2.6 | 2.2 | 2 | 3 | 2.75 | 2.65 | 2.65 | 3.2 | 3.19 | 2.2
| Mean |  | 2.5431 | 2.7683 | 2.6833 | 1.9552 |  | 2.515 | 2.3333 | 2.5917 | 2.5867 | 2.46 | 2.5433 | 2.5783 | 2.5154 | 2 | 2.5167 | 2.2667 | 2.3417 | 3 | 2.735 | 2.5817 | 2.6767 | 3.0833 | 2.8533 | 2.3
| Median |  | 2.5416 | 2.75 | 2.65 | 1.9602 |  | 2.59 | 2 | 2.58 | 2.585 | 2.465 | 2.555 | 2.59 | 2.5234 | 2 | 2.5 | 2.3 | 2.385 | 3 | 2.745 | 2.585 | 2.665 | 3.1 | 3.06 | 2.3
| Std Dev. |  | 0.023857 | 0.067355 | 0.098319 | 0.048596 |  | 0.17897 | 0.5164 | 0.031885 | 0.026583 | 0.032249 | 0.060222 | 0.049565 | 0.087727 | 0 | 0.075277 | 0.05164 | 0.20233 | 0 | 0.068337 | 0.074409 | 0.046332 | 0.075277 | 0.38908 | 0.10954
| Rel.Std.Dev. |  | 0.9381% | 2.433% | 3.664% | 2.485% |  | 7.116% | 22.13% | 1.23% | 1.028% | 1.311% | 2.368% | 1.922% | 3.488% | 0% | 2.991% | 2.278% | 8.64% | 0% | 2.499% | 2.882% | 1.731% | 2.441% | 13.64% | 4.763%
| PDM3 |  | -0.6597% | 8.137% | 4.817% | -23.62% |  | -1.758% | -8.855% | 1.236% | 1.041% | -3.907% | -0.6515% | 0.7156% | -1.741% | -21.88% | -1.693% | -11.46% | -8.529% | 17.19% | 6.835% | 0.8459% | 4.557% | 20.44% | 11.46% | -10.16%
| Z-Score (Absolute) |  | 0 | 1.82 | 1.13 | 4.76 | 0.35 | 0.23 | NA | 0.39 | 0.35 | 0.67 | 0 | 0.28 | 0.23 | NA | 0.22 | 2.24 | 1.63 | NA | 1.55 | 0.31 | 1.08 | 4.37 | 2.51 | 1.97
Results from laboratories 1.13, 1.20 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 284 (Certified Value 0.67 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.62807 | 0.6 | 0.58 | < 0.1 | < 5 | 2.98 | 1.39 | 0.83 | 0.62 | 0.63 | 1.12 | 0.52 | 0.73334 | 0.7 | 0.74 | 0.8 | 0.64 | 0.7 | 0.62 | 0.68 | 0.59 | 0.72 | 0.57 |  |  |  | 1
| 1 | 2 | 0.6723 | 0.59 | 0.58 | < 0.1 | < 5 | 0.88 | 0.54 | 0.61 | 0.55 | 0.63 | 0.55 | 0.5 | 0.65926 | 0.6 | 0.8 | 0.74 | 0.74 | 0.6 | 0.6 | 0.79 | 0.74 | 0.71 | 0.65 |  |  |  | 26
| 1 | 3 | 0.66033 | 0.78 | 0.58 | < 0.1 | < 5 | 0.71 | 0.78 | 0.58 | 0.5 | 0.62 | 0.56 | 0.54 | 0.72749 | 0.6 | 0.76 | 0.86 | 0.59 | 0.7 | 0.87 | 0.82 | 0.73 | 0.784 | 0.7 |  |  |  | 16
| 1 | 4 | 0.66128 | 0.68 | 0.62 | < 0.1 | < 5 | 1.04 | 0.63 | 0.59 | 0.54 | 0.57 | 0.56 | 0.66 | 0.74911 | 0.5 | 0.79 | 0.8 | 0.78 | 0.7 | 0.51 | 0.65 | 0.78 | 0.63 | 0.65 |  |  |  | 0.66682
| 1 | 5 | 0.73221 | 0.63 | 0.61 | < 0.1 | < 5 | 0.84 | 1.07 | 0.62 | 0.58 | 0.63 | 0.5 | 0.59 | 0.66252 | 0.6 | 0.74 | 0.84 | 0.54 | 0.6 | 0.66 | 0.71 | 0.62 | 0.62 | 0.69 |  |  |  | 20
| 1 | 6 | 0.69263 | 0.59 | 0.92 | < 0.1 | < 5 | 0.83 | 0.65 | 0.61 | 0.59 | 0.58 | 0.93 | 1.16 | 0.74126 | 0.5 | 0.77 | 0.7 | 0.87 | 0.9 | 0.65 | 0.69 | 0.74 | 0.63 | 0.58
| Mean |  | 0.67447 | 0.645 | 0.64833 |  |  | 1.2133 | 0.84333 | 0.64 | 0.56333 | 0.61 | 0.70333 | 0.66167 | 0.71216 | 0.58333 | 0.76667 | 0.79 | 0.69333 | 0.7 | 0.65167 | 0.72333 | 0.7 | 0.68233 | 0.64
| Median |  | 0.66679 | 0.615 | 0.595 |  |  | 0.86 | 0.715 | 0.61 | 0.565 | 0.625 | 0.56 | 0.565 | 0.73042 | 0.6 | 0.765 | 0.8 | 0.69 | 0.7 | 0.635 | 0.7 | 0.735 | 0.67 | 0.65
| Std Dev. |  | 0.035199 | 0.074498 | 0.13423 |  |  | 0.872 | 0.32531 | 0.094234 | 0.042269 | 0.027568 | 0.25727 | 0.25079 | 0.040394 | 0.075277 | 0.025033 | 0.060332 | 0.12485 | 0.10954 | 0.11957 | 0.066833 | 0.076158 | 0.066156 | 0.054406
| Rel.Std.Dev. |  | 5.219% | 11.55% | 20.7% |  |  | 71.87% | 38.57% | 14.72% | 7.503% | 4.519% | 36.58% | 37.9% | 5.672% | 12.9% | 3.265% | 7.637% | 18.01% | 15.65% | 18.35% | 9.24% | 10.88% | 9.696% | 8.501%
| PDM3 |  | 1.147% | -3.273% | -2.773% |  |  | 81.96% | 26.47% | -4.022% | -15.52% | -8.521% | 5.475% | -0.7732% | 6.8% | -12.52% | 14.97% | 18.47% | 3.976% | 4.975% | -2.273% | 8.475% | 4.975% | 2.326% | -4.022%
| Z-Score (Absolute) |  | 0.13 | 0.61 | 0.56 | 10.4 | 29.89 | 8.73 | 2.65 | 0.7 | 1.96 | 1.19 | 0.35 | 0.34 | 0.49 | NA | 1.39 | 1.77 | 0.18 | NA | 0.5 | 0.67 | 0.29 | 0 | 0.7
Results from laboratories 1.20 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 284 (Certified Value 1.87 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.7861 | 1.8167 | 1.8922 | 1.88 | 1.9579 | 1.87 | 1.9696 | 1.95 | 1.87 | 2.04 | 1.83 | 1.76 | 1.9 | 1.9315 | 1.74 | 1.72 | 1.82 | 1.78 | 1.79 | 1.757 | 1.8772 | 1.8504 | 1.866 | 1.872 | 1.79 | 1.76 | 1
| 1 | 2 | 1.8107 | 1.8147 | 1.8774 | 1.86 | 1.9271 | 1.86 | 1.9714 | 2.03 | 1.92 | 2 | 1.83 | 1.84 | 2.02 | 1.9344 | 1.76 | 1.76 | 1.81 | 1.78 | 1.82 | 1.856 | 1.8549 | 1.8939 | 2.069 | 1.908 | 1.79 | 1.74 | 0
| 1 | 3 | 1.8014 | 1.8688 | 1.9371 | 1.93 | 1.9287 | 1.85 | 1.9603 | 2.01 | 1.87 | 1.99 | 1.8 | 1.84 | 1.99 | 1.9144 | 1.71 | 1.8 | 1.83 | 1.75 | 1.81 | 1.779 | 1.8355 | 1.9141 | 1.91 | 1.844 | 1.79 | 1.72 | 16
| 1 | 4 | 1.7782 | 1.9078 | 1.9027 | 1.93 | 1.9603 | 1.88 | 2.0056 | 1.89 | 1.87 | 1.99 | 1.89 | 1.88 | 1.97 | 1.923 | 1.71 | 1.76 | 1.79 | 1.79 | 1.81 | 1.853 | 1.8768 | 1.9015 | 2.217 | 1.851 | 1.77 | 1.81 | 1.8676
| 1 | 5 | 1.8086 | 1.82 | 1.9074 | 1.85 | 1.8953 | 1.86 | 1.9449 | 1.95 | 1.93 | 2.03 | 1.85 | 1.83 | 1.99 | 1.918 | 1.75 | 1.8 | 1.83 | 1.76 | 1.86 | 1.855 | 1.9139 | 1.9341 | 1.948 | 1.901 | 1.82 | 1.72 | 21
| 1 | 6 | 1.8066 | 1.8345 | 1.9393 | 1.92 | 1.9377 | 1.9 | 1.9617 | 1.86 | 1.89 | 2.02 | 1.84 | 1.8 | 2.02 | 1.9198 | 1.71 | 1.77 | 1.85 | 1.78 | 1.84 | 1.848 | 1.8896 | 1.9261 | 2.047 | 1.951 | 1.79 | 1.74
| Mean |  | 1.7986 | 1.8438 | 1.9094 | 1.895 | 1.9345 | 1.87 | 1.9689 | 1.9483 | 1.8917 | 2.0117 | 1.84 | 1.825 | 1.9817 | 1.9235 | 1.73 | 1.7683 | 1.8217 | 1.7733 | 1.8217 | 1.8247 | 1.8746 | 1.9034 | 2.0095 | 1.8878 | 1.7917 | 1.7483
| Median |  | 1.804 | 1.8273 | 1.905 | 1.9 | 1.9332 | 1.865 | 1.9656 | 1.95 | 1.88 | 2.01 | 1.835 | 1.835 | 1.99 | 1.9214 | 1.725 | 1.765 | 1.825 | 1.78 | 1.815 | 1.8505 | 1.877 | 1.9078 | 1.9975 | 1.8865 | 1.79 | 1.74
| Std Dev. |  | 0.013347 | 0.03729 | 0.024617 | 0.036194 | 0.023868 | 0.017889 | 0.020275 | 0.065853 | 0.027142 | 0.02137 | 0.029665 | 0.040866 | 0.04446 | 0.0078749 | 0.022804 | 0.029944 | 0.020412 | 0.015055 | 0.024833 | 0.044527 | 0.027195 | 0.029917 | 0.12828 | 0.040246 | 0.016021 | 0.033714
| Rel.Std.Dev. |  | 0.7421% | 2.022% | 1.289% | 1.91% | 1.234% | 0.9566% | 1.03% | 3.38% | 1.435% | 1.062% | 1.612% | 2.239% | 2.244% | 0.4094% | 1.318% | 1.693% | 1.121% | 0.849% | 1.363% | 2.44% | 1.451% | 1.572% | 6.384% | 2.132% | 0.8942% | 1.928%
| PDM3 |  | -3.692% | -1.274% | 2.238% | 1.47% | 3.585% | 0.1311% | 5.428% | 4.326% | 1.291% | 7.717% | -1.475% | -2.278% | 6.11% | 2.997% | -7.365% | -5.313% | -2.457% | -5.045% | -2.457% | -2.296% | 0.3801% | 1.917% | 7.601% | 1.086% | -4.063% | -6.384%
| Z-Score (Absolute) |  | 0.98 | 0.38 | 0.49 | 0.3 | 0.82 | 0.03 | 1.28 | 1.01 | 0.26 | 1.84 | 0.43 | 0.63 | 1.45 | 0.68 | 1.88 | 1.38 | 0.67 | 1.31 | 0.67 | 0.63 | 0.03 | 0.41 | 1.82 | 0.21 | 1.07 | 1.64
| Analytical results for Cd in OREAS 284 (Certified Value 0.2 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.20776 | 0.23 | < 0.3 | 0.7342 | 0.9 | 0.28 | 0.28 | 0.18 | 0.2 | 0.21 | 0.21 | 0.22 | 0.19509 | 0.2 | 0.18 | 0.16 | 0.2 | 0.2 | 0.2 | 0.21 | 0.21 | < 1 | 0.26 | 0.17 |  |  | 1
| 1 | 2 | 0.20394 | 0.18 | < 0.3 | 0.4141 | 0.9 | 0.24 | 0.25 | 0.19 | 0.2 | 0.21 | 0.19 | 0.21 | 0.1563 | 0.2 | 0.18 | 0.15 | 0.21 | 0.2 | 0.21 | 0.2 | 0.23 | 1 | 0.24 | 0.21 |  |  | 27
| 1 | 3 | 0.21038 | 0.2 | < 0.3 | 0.6887 | 0.8 | 0.24 | 0.32 | 0.18 | 0.2 | 0.18 | 0.15 | 0.2 | 0.17276 | 0.2 | 0.19 | 0.15 | 0.23 | 0.2 | 0.2 | 0.21 | 0.2 | 1 | 0.22 | 0.2 |  |  | 16
| 1 | 4 | 0.22402 | 0.18 | < 0.3 | 0.4959 | 0.9 | 0.24 | 0.34 | 0.17 | 0.21 | 0.2 | 0.19 | 0.21 | 0.22428 | 0.2 | 0.18 | 0.13 | 0.17 | 0.2 | 0.2 | 0.18 | 0.21 | < 1 | 0.23 | 0.2 |  |  | 0.2027
| 1 | 5 | 0.20207 | 0.19 | < 0.3 | 0.4582 | 0.9 | 0.26 | 0.33 | 0.21 | 0.19 | 0.17 | 0.18 | 0.19 | 0.2171 | 0.2 | 0.19 | 0.16 | 0.2 | 0.2 | 0.2 | 0.18 | 0.23 | < 1 | 0.21 | 0.17 |  |  | 22
| 1 | 6 | < 0.2 | 0.19 | < 0.3 | 0.5957 | 0.9 | 0.24 | 0.25 | 0.19 | 0.21 | 0.2 | 0.22 | 0.2 | 0.2341 | 0.2 | 0.18 | 0.16 | 0.22 | 0.2 | 0.21 | 0.18 | 0.2 | 1 | 0.23 | 0.21
| Mean |  | 0.20963 | 0.195 |  | 0.56447 | 0.88333 | 0.25 | 0.295 | 0.18667 | 0.20167 | 0.195 | 0.19 | 0.205 | 0.19994 | 0.2 | 0.18333 | 0.15167 | 0.205 | 0.2 | 0.20333 | 0.19333 | 0.21333 | 1 | 0.23167 | 0.19333
| Median |  | 0.20776 | 0.19 |  | 0.5458 | 0.9 | 0.24 | 0.3 | 0.185 | 0.2 | 0.2 | 0.19 | 0.205 | 0.2061 | 0.2 | 0.18 | 0.155 | 0.205 | 0.2 | 0.2 | 0.19 | 0.21 | 1 | 0.23 | 0.2
| Std Dev. |  | 0.0086677 | 0.018708 |  | 0.12948 | 0.040825 | 0.016733 | 0.040373 | 0.013663 | 0.0075277 | 0.016432 | 0.024495 | 0.010488 | 0.030727 | 3.0405e-17 | 0.005164 | 0.01169 | 0.020736 | 3.0405e-17 | 0.005164 | 0.015055 | 0.013663 | 0 | 0.017224 | 0.018619
| Rel.Std.Dev. |  | 4.135% | 9.594% |  | 22.94% | 4.622% | 6.693% | 13.69% | 7.319% | 3.733% | 8.427% | 12.89% | 5.116% | 15.37% | 1.52e-14% | 2.817% | 7.708% | 10.12% | 1.52e-14% | 2.54% | 7.787% | 6.404% | 0% | 7.435% | 9.631%
| PDM3 |  | 3.419% | -3.8% |  | 178.5% | 335.8% | 23.33% | 45.53% | -7.911% | -0.511% | -3.8% | -6.267% | 1.133% | -1.363% | -1.333% | -9.555% | -25.18% | 1.133% | -1.333% | 0.3112% | -4.622% | 5.245% | 393.3% | 14.29% | -4.622%
| Z-Score (Absolute) |  | 0.55 | 0.22 | 4.25 | 32.86 | NA | 4.7 | 8.73 | 0.97 | 0.38 | 0.22 | 0.67 | 0.67 | 0.22 | NA | 1.27 | 4.1 | 0.67 | NA | 0.53 | 0.37 | 1.42 | NA | 3.06 | 0.37
Results from laboratories 1.11, 1.20 and 1.24 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 284 (Certified Value 79 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 77.2971 | 80.92 | 78.2 | 85.3289 | 48.2 | 70.85 | 72.4 | 72.7 | 76.3 | 71.5 | 82.5 | 82 | 73.52 | 81.3 | 78 | 75.46 | 83 | 78.3 | 78 | 55.4 | 69.03 | [5 empty] | 1
| 1 | 2 | 76.4251 | 82.74 | 78.5 | 86.6834 | 49.7 | 66.28 | 79.2 | 78.7 | 71.3 | 76.6 | 79 | 96 | 72.54 | 86.9 | 81 | 80.34 | 81 | 78.23 | 79 | 53.4 | 64.36 | [5 empty] | 28
| 1 | 3 | 77.5679 | 81.52 | 80 | 86.7233 | 49.9 | 73.54 | 75.1 | 78.1 | 75 | 76.1 | 78.3 | 95 | 73.78 | 86.8 | 76 | 78.19 | 82 | 78.14 | 77 | 56.5 | 63.01 | [5 empty] | 16
| 1 | 4 | 79.0859 | 80.05 | 76.7 | 80.7039 | 50.9 | 68.5 | 76.2 | 83.4 | 78.1 | 75.9 | 82.7 | 93 | 74.7 | 87 | 78 | 79.91 | 81 | 79.53 | 76 | 51.7 | 64.68 | [5 empty] | 78.9981
| 1 | 5 | 76.7178 | 81.95 | 74.4 | 78.7871 | 50.4 | 72.49 | 81.3 | 78.4 | 77.3 | 71 | 79.3 | 84 | 74.09 | 88.6 | 81 | 81.59 | 82 | 79.34 | 78 | 50.6 | 64.93 | [5 empty] | 23
| 1 | 6 | 76.9118 | 82.27 | 80.1 | 80.2896 | 49.7 | 68.67 | 77.5 | 84 | 76.3 | 73.8 | 80.1 | 91 | 71.81 | 89.1 | 76 | 80.98 | 83 | 78.66 | 78 | 53.1 | 65.93
| Mean |  | 77.3343 | 81.575 | 77.9833 | 83.086 | 49.8 | 70.055 | 76.95 | 79.2167 | 75.7167 | 74.15 | 80.3167 | 90.1667 | 73.4067 | 86.6167 | 78.3333 | 79.4117 | 82 | 78.7 | 77.6667 | 53.45 | 65.3233
| Median |  | 77.1044 | 81.735 | 78.35 | 83.0164 | 49.8 | 69.76 | 76.85 | 78.55 | 76.3 | 74.85 | 79.7 | 92 | 73.65 | 86.95 | 78 | 80.125 | 82 | 78.48 | 78 | 53.25 | 64.805
| Std Dev. |  | 0.94966 | 0.9735 | 2.1609 | 3.5546 | 0.91214 | 2.7311 | 3.1297 | 4.1238 | 2.4037 | 2.4468 | 1.8616 | 5.8452 | 1.0576 | 2.7795 | 2.2509 | 2.2549 | 0.89443 | 0.5991 | 1.0328 | 2.2098 | 2.0472
| Rel.Std.Dev. |  | 1.228% | 1.193% | 2.771% | 4.278% | 1.832% | 3.898% | 4.067% | 5.206% | 3.175% | 3.3% | 2.318% | 6.483% | 1.441% | 3.209% | 2.874% | 2.84% | 1.091% | 0.7612% | 1.33% | 4.134% | 3.134%
| PDM3 |  | -2.106% | 3.262% | -1.285% | 5.175% | -36.96% | -11.32% | -2.593% | 0.2767% | -4.154% | -6.137% | 1.669% | 14.14% | -7.078% | 9.644% | -0.8415% | 0.5236% | 3.8% | -0.3773% | -1.685% | -32.34% | -17.31%
| Z-Score (Absolute) |  | 0.12 | 0.67 | 0 | 0.96 | 5.29 | 1.49 | 0.19 | 0.23 | 0.43 | 0.72 | 0.44 | 2.29 | 0.86 | 1.62 | 0.07 | 0.27 | 0.75 | 0.13 | 0.06 | 4.61 | 2.38
| Analytical results for Co in OREAS 284 (Certified Value 12.3 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 18.94 | 11.841 | 12.4 | 13 | 23.9982 | 12 | 12.7 | 10.3 | 11.9 | 11.8 | 11.8 | 12 | 13 | 12.3325 | 11.7 | 11 | 12 | 12.8 | 13.1 | 12.1 | 12.7 | 12.3 | 14.02 | 12 | 12.1 | 11.9 | 1
| 1 | 2 | 19.36 | 11.9993 | 12.2 | 14 | 13.2175 | 12 | 12.9 | 11.9 | 12.2 | 11.7 | 10.8 | 12.6 | 14.7 | 12.4263 | 11.8 | 11.1 | 12 | 12.6 | 12.3 | 12.7 | 12.5 | 12.6 | 14.79 | 13 | 12.2 | 12 | 16
| 1 | 3 | 19.31 | 12.2815 | 12.4 | 13 | 11.485 | 12 | 13.2 | 11.6 | 11.6 | 11.3 | 11.4 | 12.6 | 13.2 | 12.1741 | 11.6 | 11.4 | 12 | 12.7 | 12.7 | 12.2 | 12.4 | 12.6 | 13.58 | 12 | 11.8 | 11.8 | 16
| 1 | 4 | 18.98 | 12.3948 | 12.4 | 13 | 15.6236 | 12 | 13.5 | 11.3 | 11.9 | 12.3 | 11.6 | 12.4 | 13.9 | 12.4464 | 11.1 | 11.4 | 12 | 12.8 | 12.4 | 12.9 | 12.5 | 12.5 | 13.24 | 15 | 12.3 | 12.1 | 12.2823
| 1 | 5 | 19 | 11.9824 | 12.2 | 13 | 12.2404 | 12 | 13.1 | 10.8 | 12.6 | 11.8 | 11.8 | 12 | 13.2 | 12.1888 | 10.9 | 11.6 | 12 | 12.8 | 12.8 | 13 | 12.9 | 12.8 | 13.3 | 13 | 12.3 | 12 | 24
| 1 | 6 | 19.22 | 12.0797 | 12.3 | 13 | 11.3924 | 12 | 13.1 | 10.8 | 12.2 | 12.8 | 11.4 | 12.6 | 13.8 | 12.434 | 11.6 | 11.6 | 12 | 13 | 12.4 | 12.5 | 12.5 | 12.7 | 12.67 | 12 | 12.7 | 11.9
| Mean |  | 19.135 | 12.0965 | 12.3167 | 13.1667 | 14.6595 | 12 | 13.0833 | 11.1167 | 12.0667 | 11.95 | 11.4667 | 12.3667 | 13.6333 | 12.3337 | 11.45 | 11.35 | 12 | 12.7833 | 12.6167 | 12.5667 | 12.5833 | 12.5833 | 13.6 | 12.8333 | 12.2333 | 11.95
| Median |  | 19.11 | 12.0395 | 12.35 | 13 | 12.7289 | 12 | 13.1 | 11.05 | 12.05 | 11.8 | 11.5 | 12.5 | 13.5 | 12.3794 | 11.6 | 11.4 | 12 | 12.8 | 12.55 | 12.6 | 12.5 | 12.6 | 13.44 | 12.5 | 12.25 | 11.95
| Std Dev. |  | 0.18371 | 0.20555 | 0.098319 | 0.40825 | 4.8337 | 0 | 0.27142 | 0.59133 | 0.34448 | 0.5244 | 0.37238 | 0.29439 | 0.63456 | 0.12475 | 0.36194 | 0.251 | 0 | 0.13292 | 0.30605 | 0.36697 | 0.18348 | 0.17224 | 0.73176 | 1.169 | 0.29439 | 0.10488
| Rel.Std.Dev. |  | 0.9601% | 1.699% | 0.7983% | 3.101% | 32.97% | 0% | 2.075% | 5.319% | 2.855% | 4.388% | 3.247% | 2.381% | 4.654% | 1.011% | 3.161% | 2.211% | 0% | 1.04% | 2.426% | 2.92% | 1.458% | 1.369% | 5.381% | 9.109% | 2.406% | 0.8777%
| PDM3 |  | 55.79% | -1.513% | 0.2794% | 7.2% | 19.35% | -2.299% | 6.521% | -9.491% | -1.756% | -2.706% | -6.641% | 0.6865% | 11% | 0.4178% | -6.777% | -7.591% | -2.299% | 4.079% | 2.722% | 2.315% | 2.451% | 2.451% | 10.73% | 4.486% | -0.3991% | -2.706%
| Z-Score (Absolute) |  | 11.43 | 0.43 | 0.06 | NA | 3.89 | NA | 1.24 | 2.08 | 0.48 | 0.67 | 1.49 | 0.03 | 2.16 | 0.03 | 1.52 | 1.69 | NA | 0.73 | 0.45 | 0.36 | 0.39 | 0.39 | 2.11 | NA | 0.2 | 0.67
Results from laboratories 1.09, 1.11, 1.22 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 284 (Certified Value 47.7 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 45.4065 | 51 | 56 | 65.9894 | 51 | 48.8 | 36 | 48 | 48 | 46 | 45 | 46 | 46.9635 | 50 | 53 | 54 | 45.4 | 45 | 31 | 49 | 50 | 51.37 | 50 | 48 | 26 |  | 1
| 1 | 2 | 45.2741 | 50 | 51 | 45.1882 | 49 | 49.1 | 33 | 47 | 47 | 46 | 48 | 48 | 40.0443 | 44 | 55 | 55 | 44.5 | 41 | 31 | 45 | 42 |  | 45 | 49 | 28 |  | 17
| 1 | 3 | 45.4065 | 48 | 48 | 46.624 | 48 | 48.7 | 36 | 46 | 46 | 44 | 48 | 49 | 37.5698 | 46 | 57 | 55 | 41.7 | 45 | 33 | 45 | 48 | 51.79 | 46 | 48 | 27 |  | 16
| 1 | 4 | 47.4757 | 50 | 54 | 45.9337 | 50 | 49.8 | 33 | 47 | 46 | 43 | 49 | 49 | 38.413 | 43 | 55 | 55 | 44.3 | 44 | 36 | 46 | 49 | 55.27 | 42 | 48 | 27 |  | 47.7435
| 1 | 5 | 45.7141 | 45 | 55 | 46.7951 | 52 | 48.3 | 36 | 47 | 47 | 43 | 47 | 49 | 41.4008 | 46 | 55 | 56 | 40.4 | 46 | 37 | 46 | 42 | 49.56 | 40 | 49 | 28 |  | 25
| 1 | 6 | 45.8934 | 51 | 49 | 44.1942 | 51 | 48.9 | 34 | 47 | 47 | 43 | 47 | 48 | 46.3718 | 48 | 55 | 56 | 45.2 | 45 | 38 | 45 | 53 | 52.15 | 43 | 49 | 28
| Mean |  | 45.8617 | 49.1667 | 52.1667 | 49.1208 | 50.1667 | 48.9333 | 34.6667 | 47 | 46.8333 | 44.1667 | 47.3333 | 48.1667 | 41.7938 | 46.1667 | 55 | 55.1667 | 43.5833 | 44.3333 | 34.3333 | 46 | 47.3333 | 52.028 | 44.3333 | 48.5 | 27.3333
| Median |  | 45.5603 | 50 | 52.5 | 46.2788 | 50.5 | 48.85 | 35 | 47 | 47 | 43.5 | 47.5 | 48.5 | 40.7225 | 46 | 55 | 55 | 44.4 | 45 | 34.5 | 45.5 | 48.5 | 51.79 | 44 | 48.5 | 27.5
| Std Dev. |  | 0.82309 | 2.3166 | 3.3116 | 8.3197 | 1.472 | 0.50067 | 1.5055 | 0.63246 | 0.75277 | 1.472 | 1.3663 | 1.169 | 4.0043 | 2.5626 | 1.2649 | 0.75277 | 2.0469 | 1.7512 | 3.0768 | 1.5492 | 4.4572 | 2.068 | 3.5024 | 0.54772 | 0.8165
| Rel.Std.Dev. |  | 1.795% | 4.712% | 6.348% | 16.94% | 2.934% | 1.023% | 4.343% | 1.346% | 1.607% | 3.333% | 2.886% | 2.427% | 9.581% | 5.551% | 2.3% | 1.365% | 4.696% | 3.95% | 8.962% | 3.368% | 9.417% | 3.975% | 7.9% | 1.129% | 2.987%
| PDM3 |  | -3.942% | 2.981% | 9.264% | 2.885% | 5.075% | 2.492% | -27.39% | -1.557% | -1.906% | -7.492% | -0.8592% | 0.8863% | -12.46% | -3.303% | 15.2% | 15.55% | -8.714% | -7.143% | -28.09% | -3.652% | -0.8592% | 8.974% | -7.143% | 1.584% | -42.75%
| Z-Score (Absolute) |  | 0.29 | 0.55 | 1.31 | 0.54 | 0.8 | 0.49 | 3.12 | 0 | 0.04 | 0.72 | 0.08 | 0.3 | 1.32 | 0.21 | 2.02 | 2.07 | 0.86 | 0.67 | 3.2 | 0.25 | 0.08 | 1.27 | 0.67 | 0.38 | 4.97
| Analytical results for Cs in OREAS 284 (Certified Value 11.4 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 10.9167 | 11.63 | 10.9 | 12.6 | 10 | 11.5 | 12.6 | 11.2 | 10.3 | 12.05 | 11.684 | 11.7 | 10.9 | 10.9 | 11.9 | 11.01 | 11.81 | 11.08 | 10.8 | 10 | [6 empty] | 1
| 1 | 2 | 10.8883 | 11.71 | 10.8 | 12.7 | 11.3 | 11.15 | 12.45 | 10.5 | 10.8 | 11.9 | 12.0115 | 11.9 | 11 | 10.7 | 11.6 | 11.74 | 11.77 | 11.53 | 10.7 | 11 | [6 empty] | 31
| 1 | 3 | 11.2909 | 11.82 | 11.2 | 12.9 | 10.9 | 11.65 | 12.2 | 10.95 | 10.75 | 11.5 | 11.7276 | 11.4 | 11.2 | 10.8 | 11.9 | 11.04 | 11.63 | 11.46 | 10.7 | 10 | [6 empty] | 16
| 1 | 4 | 11.4757 | 11.65 | 11 | 13.2 | 10.7 | 11.65 | 11.65 | 11.2 | 10.7 | 12.35 | 11.8628 | 12.3 | 11.1 | 10.6 | 12.1 | 11.54 | 11.96 | 11.31 | 11.3 | 11 | [6 empty] | 11.3727
| 1 | 5 | 11.1593 | 11.72 | 10.7 | 13 | 9.7 | 12.15 | 12.4 | 11.2 | 10.25 | 11.45 | 11.7922 | 11 | 11.5 | 10.6 | 12.3 | 11.93 | 12.1 | 11.47 | 11.1 | 11 | [6 empty] | 26
| 1 | 6 | 11.3224 | 11.74 | 10.8 | 13.3 | 9.5 | 11.95 | 11.4 | 11.05 | 10.8 | 12 | 12.004 | 11.3 | 11.3 | 10.6 | 12 | 11.48 | 11.73 | 11.45 | 11.4 | 11
| Mean |  | 11.1756 | 11.7117 | 10.9 | 12.95 | 10.35 | 11.675 | 12.1167 | 11.0167 | 10.6 | 11.875 | 11.847 | 11.6 | 11.1667 | 10.7 | 11.9667 | 11.4567 | 11.8333 | 11.3833 | 11 | 10.6667
| Median |  | 11.2251 | 11.715 | 10.85 | 12.95 | 10.35 | 11.65 | 12.3 | 11.125 | 10.725 | 11.95 | 11.8275 | 11.55 | 11.15 | 10.65 | 11.95 | 11.51 | 11.79 | 11.455 | 10.95 | 11
| Std Dev. |  | 0.23443 | 0.067946 | 0.17889 | 0.27386 | 0.72042 | 0.34893 | 0.48236 | 0.27325 | 0.25495 | 0.3446 | 0.13841 | 0.46476 | 0.21602 | 0.12649 | 0.23381 | 0.37012 | 0.16955 | 0.16537 | 0.30984 | 0.5164
| Rel.Std.Dev. |  | 2.098% | 0.5802% | 1.641% | 2.115% | 6.961% | 2.989% | 3.981% | 2.48% | 2.405% | 2.902% | 1.168% | 4.007% | 1.935% | 1.182% | 1.954% | 3.231% | 1.433% | 1.453% | 2.817% | 4.841%
| PDM3 |  | -1.733% | 2.981% | -4.156% | 13.87% | -8.992% | 2.658% | 6.542% | -3.13% | -6.794% | 4.417% | 4.171% | 1.999% | -1.811% | -5.915% | 5.223% | 0.7385% | 4.051% | 0.09368% | -3.277% | -6.208%
| Z-Score (Absolute) |  | 0.45 | 0.41 | 0.9 | 2.41 | 1.78 | 0.35 | 1.06 | 0.71 | 1.38 | 0.67 | 0.63 | 0.23 | 0.47 | 1.22 | 0.82 | 0 | 0.61 | 0.12 | 0.74 | NA
Results from laboratory 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 284 (Certified Value 27.6 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 32.14 | 27.6045 | 28.3 | 27 | 38.2229 | 32 | 39.5 | 29.2 | 27.5 | 25.7 | 26 | 24.3 | 27.6 | 26.1986 | 24.4 | 29.2 | 24 | 27.3 | 28.1 | 26.6 | 26 | 26.9 | 33.82 | 46 | 33.1 | 26.3 | 1
| 1 | 2 | 31.72 | 27.9216 | 27.8 | 28 | 32.3798 | 31 | 29.3 | 29.2 | 27.9 | 25.5 | 25.5 | 26.2 | 28.7 | 25.5272 | 22.7 | 28.8 | 25 | 27.1 | 25.8 | 28.4 | 26 | 28.5 | 32.27 | 43 | 33.1 | 24.9 | 18
| 1 | 3 | 32.16 | 27.9278 | 28.1 | 30 | 30.1538 | 34 | 31.1 | 29.4 | 25.8 | 24.6 | 25.8 | 26.2 | 26.1 | 27.1479 | 23.5 | 29.4 | 25 | 26.9 | 26.6 | 27.6 | 26 | 27.9 | 26.78 | 36 | 33.2 | 43 | 16
| 1 | 4 | 31.52 | 28.8197 | 28.1 | 28 | 30.8589 | 33 | 29.8 | 27 | 26.3 | 28.6 | 26.2 | 25.8 | 27.6 | 25.6604 | 22.5 | 29 | 26 | 27 | 26.5 | 29.1 | 26 | 27.4 | 28.65 | 37 | 33.8 | 24.8 | 27.5974
| 1 | 5 | 32.81 | 27.4691 | 27.9 | 28 | 28.4924 | 30 | 30.4 | 28.6 | 27.7 | 25.2 | 26.2 | 25.6 | 25.6 | 26.582 | 21.5 | 29.4 | 25 | 26.7 | 28.8 | 28.1 | 27 | 27.1 | 27.24 | 36 | 33.2 | 23.7 | 27
| 1 | 6 | 32.93 | 27.1584 | 28.4 | 28 | 30.0239 | 30 | 29 | 27.4 | 26.8 | 30.3 | 25.3 | 25.9 | 27.2 | 27.2839 | 23.5 | 29.1 | 25 | 26.7 | 26.6 | 28.6 | 27 | 28.5 | 27.15 | 37 | 34.1 | 24.1
| Mean |  | 32.2133 | 27.8169 | 28.1 | 28.1667 | 31.6886 | 31.6667 | 31.5167 | 28.4667 | 27 | 26.65 | 25.8333 | 25.6667 | 27.1333 | 26.4 | 23.0167 | 29.15 | 25 | 26.95 | 27.0667 | 28.0667 | 26.3333 | 27.7167 | 29.3183 | 39.1667 | 33.4167 | 27.8
| Median |  | 32.15 | 27.7631 | 28.1 | 28 | 30.5064 | 31.5 | 30.1 | 28.9 | 27.15 | 25.6 | 25.9 | 25.85 | 27.4 | 26.3903 | 23.1 | 29.15 | 25 | 26.95 | 26.6 | 28.25 | 26 | 27.65 | 27.945 | 37 | 33.2 | 24.85
| Std Dev. |  | 0.56603 | 0.57069 | 0.22804 | 0.98319 | 3.4411 | 1.633 | 3.9837 | 1.025 | 0.83905 | 2.2652 | 0.37238 | 0.70899 | 1.1237 | 0.73798 | 1.0048 | 0.23452 | 0.63246 | 0.23452 | 1.1343 | 0.8756 | 0.5164 | 0.69402 | 2.9963 | 4.2622 | 0.42622 | 7.4993
| Rel.Std.Dev. |  | 1.757% | 2.052% | 0.8115% | 3.491% | 10.86% | 5.157% | 12.64% | 3.601% | 3.108% | 8.5% | 1.441% | 2.762% | 4.141% | 2.795% | 4.366% | 0.8045% | 2.53% | 0.8702% | 4.191% | 3.12% | 1.961% | 2.504% | 10.22% | 10.88% | 1.275% | 26.98%
| PDM3 |  | 16.73% | 0.7953% | 1.821% | 2.063% | 14.82% | 14.75% | 14.2% | 3.15% | -2.165% | -3.433% | -6.392% | -6.996% | -1.681% | -4.339% | -16.6% | 5.626% | -9.412% | -2.346% | -1.923% | 1.7% | -4.58% | 0.4323% | 6.236% | 41.92% | 21.09% | 0.7342%
| Z-Score (Absolute) |  | 2.16 | 0 | 0.14 | 0.18 | 1.9 | 1.89 | 1.82 | 0.32 | 0.4 | 0.57 | 0.97 | 1.05 | 0.33 | 0.69 | 2.35 | 0.66 | 1.38 | 0.42 | 0.36 | 0.13 | 0.72 | 0.05 | 0.74 | 5.57 | 2.75 | 0
| Analytical results for Dy in OREAS 284 (Certified Value 3.94 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.02: 1.03; 1.03: 1.04; 1.04: 1.09; 1.09: 1.10; 1.10: 1.13; 1.11: 1.26; 1.12: 1.29 | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS | 2
3
| 1 | 1 | 1.02: 3.8102; 1.03: 3.9; 1.04: 4.2; 1.09: 3.7677; 1.10: 3.3; 1.11: 4.07; 1.12: 4 | 1
| 1 | 2 | 1.02: 3.7639; 1.03: 4.09; 1.04: 4; 1.09: 3.7405; 1.10: 4.1; 1.11: 4.1; 1.12: 4.1 | 33
| 1 | 3 | 1.02: 3.8868; 1.03: 3.93; 1.04: 4.3; 1.09: 3.7978; 1.10: 3.7; 1.11: 3.96; 1.12: 4 | 16
| 1 | 4 | 1.02: 3.9086; 1.03: 4.13; 1.04: 3.9; 1.09: 3.6403; 1.10: 3.6; 1.11: 4.04; 1.12: 4 | 3.9404
| 1 | 5 | 1.02: 3.8398; 1.03: 3.95; 1.04: 3.7; 1.09: 3.4829; 1.10: 3.5; 1.11: 4.16; 1.12: 4.2 | 28
| 1 | 6 | 1.02: 3.8197; 1.03: 3.91; 1.04: 4; 1.09: 3.4966; 1.10: 3.4; 1.11: 4.16; 1.12: 4.1
| Mean |  | 1.02: 3.8382; 1.03: 3.985; 1.04: 4.0167; 1.09: 3.6543; 1.10: 3.6; 1.11: 4.0817; 1.12: 4.0667
| Median |  | 1.02: 3.8298; 1.03: 3.94; 1.04: 4; 1.09: 3.6904; 1.10: 3.55; 1.11: 4.085; 1.12: 4.05
| Std Dev. |  | 1.02: 0.052856; 1.03: 0.099146; 1.04: 0.2137; 1.09: 0.13808; 1.10: 0.28284; 1.11: 0.076529; 1.12: 0.08165
| Rel.Std.Dev. |  | 1.02: 1.377%; 1.03: 2.488%; 1.04: 5.32%; 1.09: 3.778%; 1.10: 7.857%; 1.11: 1.875%; 1.12: 2.008%
| PDM3 |  | 1.02: -2.595%; 1.03: 1.132%; 1.04: 1.935%; 1.09: -7.261%; 1.10: -8.639%; 1.11: 3.585%; 1.12: 3.204%
| Z-Score (Absolute) |  | 1.02: 1.02; 1.03: 0; 1.04: 0.22; 1.09: 2.31; 1.10: 2.69; 1.11: 0.67; 1.12: 0.57
| Analytical results for Er in OREAS 284 (Certified Value 1.42 ppm) | [27 empty] | Certified
Er | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.02: 1.03; 1.03: 1.04; 1.04: 1.09; 1.09: 1.10; 1.10: 1.13; 1.11: 1.26; 1.12: 1.29 | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS | 2
3
| 1 | 1 | 1.02: 1.5073; 1.03: 1.42; 1.04: 1.8; 1.09: 1.4153; 1.10: 1.3; 1.11: 1.44; 1.12: 1.4 | 1
| 1 | 2 | 1.02: 1.4566; 1.03: 1.43; 1.04: 1.7; 1.09: 1.3732; 1.10: 1.6; 1.11: 1.45; 1.12: 1.4 | 34
| 1 | 3 | 1.02: 1.5189; 1.03: 1.4; 1.04: 1.9; 1.09: 1.3976; 1.10: 1.4; 1.11: 1.43; 1.12: 1.3 | 16
| 1 | 4 | 1.02: 1.5206; 1.03: 1.42; 1.04: 1.7; 1.09: 1.4145; 1.10: 1.4; 1.11: 1.47; 1.12: 1.5 | 1.4199
| 1 | 5 | 1.02: 1.4917; 1.03: 1.43; 1.04: 1.6; 1.09: 1.2271; 1.10: 1.3; 1.11: 1.53; 1.12: 1.5 | 29
| 1 | 6 | 1.02: 1.4893; 1.03: 1.4; 1.04: 1.7; 1.09: 1.2855; 1.10: 1.3; 1.11: 1.5; 1.12: 1.3
| Mean |  | 1.02: 1.4974; 1.03: 1.4167; 1.04: 1.7333; 1.09: 1.3522; 1.10: 1.3833; 1.11: 1.47; 1.12: 1.4
| Median |  | 1.02: 1.4995; 1.03: 1.42; 1.04: 1.7; 1.09: 1.3854; 1.10: 1.35; 1.11: 1.46; 1.12: 1.4
| Std Dev. |  | 1.02: 0.023921; 1.03: 0.013663; 1.04: 0.10328; 1.09: 0.078056; 1.10: 0.1169; 1.11: 0.038471; 1.12: 0.089443
| Rel.Std.Dev. |  | 1.02: 1.598%; 1.03: 0.9644%; 1.04: 5.958%; 1.09: 5.773%; 1.10: 8.451%; 1.11: 2.617%; 1.12: 6.389%
| PDM3 |  | 1.02: 5.455%; 1.03: -0.2299%; 1.04: 22.07%; 1.09: -4.77%; 1.10: -2.577%; 1.11: 3.526%; 1.12: -1.404%
| Z-Score (Absolute) |  | 1.02: 1.02; 1.03: 0; 1.04: 4; 1.09: 0.82; 1.10: 0.42; 1.11: 0.67; 1.12: 0.21
| Analytical results for Eu in OREAS 284 (Certified Value 1.48 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.02: 1.03; 1.03: 1.04; 1.04: 1.09; 1.09: 1.10; 1.10: 1.13; 1.11: 1.26; 1.12: 1.29 | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS | 2
3
| 1 | 1 | 1.02: 1.3986; 1.03: 1.46; 1.04: 1.38; 1.09: 1.5412; 1.10: 1.3; 1.11: 1.57; 1.12: 1.7 | 1
| 1 | 2 | 1.02: 1.4363; 1.03: 1.46; 1.04: 1.44; 1.09: 1.5353; 1.10: 1.6; 1.11: 1.51; 1.12: 1.6 | 35
| 1 | 3 | 1.02: 1.4884; 1.03: 1.43; 1.04: 1.49; 1.09: 1.5691; 1.10: 1.4; 1.11: 1.45; 1.12: 1.5 | 16
| 1 | 4 | 1.02: 1.4732; 1.03: 1.43; 1.04: 1.39; 1.09: 1.4818; 1.10: 1.5; 1.11: 1.54; 1.12: 1.6 | 1.4774
| 1 | 5 | 1.02: 1.4386; 1.03: 1.48; 1.04: 1.29; 1.09: 1.3912; 1.10: 1.4; 1.11: 1.54; 1.12: 1.6 | 30
| 1 | 6 | 1.02: 1.4734; 1.03: 1.49; 1.04: 1.45; 1.09: 1.4245; 1.10: 1.3; 1.11: 1.5; 1.12: 1.6
| Mean |  | 1.02: 1.4514; 1.03: 1.4583; 1.04: 1.4067; 1.09: 1.4905; 1.10: 1.4167; 1.11: 1.5183; 1.12: 1.6
| Median |  | 1.02: 1.4559; 1.03: 1.46; 1.04: 1.415; 1.09: 1.5086; 1.10: 1.4; 1.11: 1.525; 1.12: 1.6
| Std Dev. |  | 1.02: 0.03318; 1.03: 0.024833; 1.04: 0.070048; 1.09: 0.070778; 1.10: 0.1169; 1.11: 0.041673; 1.12: 0.063246
| Rel.Std.Dev. |  | 1.02: 2.286%; 1.03: 1.703%; 1.04: 4.98%; 1.09: 4.749%; 1.10: 8.252%; 1.11: 2.745%; 1.12: 3.953%
| PDM3 |  | 1.02: -1.76%; 1.03: -1.292%; 1.04: -4.789%; 1.09: 0.8865%; 1.10: -4.112%; 1.11: 2.769%; 1.12: 8.297%
| Z-Score (Absolute) |  | 1.02: 0.11; 1.03: 0; 1.04: 0.84; 1.09: 0.52; 1.10: 0.67; 1.11: 0.97; 1.12: 2.29
| Analytical results for Fe in OREAS 284 (Certified Value 3.15 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.132 | 3.0995 | 3.18 | 3.15 | 3.1021 | 3.3 | 3.1494 | 3.15 | 3.09 | 3.29 | 3.02 | 2.95 | 2.89 | 3.2954 | 3.01 | 3 | 3.14 | 3.17 | 3.4 | 3.07 | 3.14 | 3.11 | 2.899 | 3.259 | 3.09 | 3.08 | 1
| 1 | 2 | 3.1721 | 3.1136 | 3.15 | 3.14 | 3.0858 | 3.3 | 3.1391 | 3.11 | 3.14 | 3.23 | 2.99 | 3.1 | 3.1 | 3.2991 | 3.02 | 3.04 | 3.21 | 3.33 | 3.05 | 3.24 | 3.14 | 3.17 |  | 3.301 | 3.13 | 3.02 | 11
| 1 | 3 | 3.1511 | 3.1953 | 3.26 | 3.22 | 3.0895 | 3.29 | 3.1218 | 3.24 | 3.04 | 3.18 | 2.98 | 3.1 | 3.04 | 3.2591 | 2.98 | 3.06 | 3.3 | 3.22 | 3.17 | 3.11 | 3.13 | 3.17 | 2.849 | 3.21 | 3.1 | 2.99 | 16
| 1 | 4 | 3.1523 | 3.2156 | 3.21 | 3.26 | 3.128 | 3.35 | 3.199 | 3.05 | 3.1 | 3.21 | 3.09 | 3.2 | 3.03 | 3.2558 | 2.96 | 3 | 3.16 | 3.27 | 3.1 | 3.25 | 3.12 | 3.19 | 2.921 | 3.203 | 3.23 | 3.11 | 3.1542
| 1 | 5 | 3.1878 | 3.1063 | 3.19 | 3.22 | 3.0359 | 3.31 | 3.1039 | 3.16 | 3.13 | 3.27 | 3.04 | 3.1 | 3.07 | 3.2346 | 2.8 | 3.06 | 3.31 | 3.12 | 3.21 | 3.26 | 3.22 | 3.19 | 2.764 | 3.308 | 3.2 | 2.99 | 31
| 1 | 6 | 3.1846 | 3.1384 | 3.27 | 3.22 | 3.0927 | 3.36 | 3.1451 | 3.01 | 3.11 | 3.24 | 2.99 | 3.08 | 3.12 | 3.2772 | 2.97 | 3.01 | 3.32 | 3.21 | 3.15 | 3.25 | 3.13 | 3.19 | 2.802 | 3.35 | 3.3 | 3.04
| Mean |  | 3.1633 | 3.1448 | 3.21 | 3.2017 | 3.089 | 3.3183 | 3.1431 | 3.12 | 3.1017 | 3.2367 | 3.0183 | 3.0883 | 3.0417 | 3.2702 | 2.9567 | 3.0283 | 3.24 | 3.22 | 3.18 | 3.1967 | 3.1467 | 3.17 | 2.847 | 3.2718 | 3.175 | 3.0383
| Median |  | 3.1622 | 3.126 | 3.2 | 3.22 | 3.0911 | 3.305 | 3.1421 | 3.13 | 3.105 | 3.235 | 3.005 | 3.1 | 3.055 | 3.2682 | 2.975 | 3.025 | 3.255 | 3.215 | 3.16 | 3.245 | 3.135 | 3.18 | 2.849 | 3.28 | 3.165 | 3.03
| Std Dev. |  | 0.021797 | 0.049225 | 0.046904 | 0.046655 | 0.030121 | 0.029269 | 0.032171 | 0.082462 | 0.035449 | 0.039833 | 0.041673 | 0.080104 | 0.081833 | 0.02497 | 0.080166 | 0.028577 | 0.08025 | 0.073756 | 0.12133 | 0.083825 | 0.036697 | 0.030984 | 0.06538 | 0.058301 | 0.082644 | 0.048751
| Rel.Std.Dev. |  | 0.689% | 1.565% | 1.461% | 1.457% | 0.9751% | 0.882% | 1.024% | 2.643% | 1.143% | 1.231% | 1.381% | 2.594% | 2.69% | 0.7636% | 2.711% | 0.9437% | 2.477% | 2.291% | 3.815% | 2.622% | 1.166% | 0.9774% | 2.296% | 1.782% | 2.603% | 1.605%
| PDM3 |  | 0.289% | -0.2995% | 1.769% | 1.505% | -2.067% | 5.203% | -0.3538% | -1.085% | -1.666% | 2.614% | -4.308% | -2.089% | -3.568% | 3.677% | -6.263% | -3.991% | 2.72% | 2.086% | 0.8176% | 1.346% | -0.2392% | 0.5006% | -9.74% | 3.729% | 0.6591% | -3.674%
| Z-Score (Absolute) |  | 0.09 | 0.11 | 0.57 | 0.48 | 0.68 | 1.68 | 0.12 | 0.36 | 0.55 | 0.84 | 1.41 | 0.69 | 1.17 | 1.19 | 2.04 | 1.3 | 0.88 | 0.67 | 0.26 | 0.43 | 0.09 | 0.15 | 3.17 | 1.21 | 0.21 | 1.2
| Analytical results for Ga in OREAS 284 (Certified Value 21 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 20.2288 | 21.6 | 22.9 | 21.9366 | 21 | 21.27 | 17.95 | 19.1 | 20.2 | 20.7 | 19.85 | 21.5 | 22.1072 | 22 | 18.6 | 20 | 23.3 | 20.11 | 21.22 | 20.42 | 20.45 | 20 | 62.1 | 18.9 |  |  | 1
| 1 | 2 | 20.1813 | 21.44 | 21.4 | 22.4563 | 22 | 22.41 | 18.99 | 19.75 | 19.3 | 19.35 | 20.4 | 21.7 | 21.8023 | 23 | 19.1 | 19.7 | 22.5 | 21.46 | 21.1 | 20.86 | 21.96 | 21 | 62.7 | 19.3 |  |  | 36
| 1 | 3 | 20.9836 | 21.78 | 22.2 | 22.8065 | 21 | 23.14 | 19.69 | 18.85 | 19.25 | 19.9 | 20.5 | 21.8 | 22.0118 | 23 | 19.3 | 19.9 | 23.2 | 20.54 | 21.18 | 21.54 | 20.87 | 20 | 61.5 | 19.3 |  |  | 16
| 1 | 4 | 21.4172 | 21.54 | 24.4 | 21.5879 | 22 | 23.89 | 19.74 | 19.1 | 20.2 | 20.4 | 20.8 | 22.9 | 21.9462 | 23 | 19.2 | 20.4 | 23.1 | 21.49 | 20.98 | 21.76 | 21.23 | 20 | 64.5 | 19.5 |  |  | 21.0112
| 1 | 5 | 21.0398 | 21.61 | 25 | 21.0009 | 21 | 23.57 | 18.36 | 20.1 | 19.7 | 20.1 | 19.65 | 21.8 | 22.3722 | 21 | 19.6 | 20.1 | 23.6 | 22.17 | 21.68 | 21.64 | 19.88 | 20 | 64.2 | 19.3 |  |  | 32
| 1 | 6 | 20.3875 | 21.95 | 23 | 21.3469 | 22 | 24.13 | 18.48 | 19.55 | 21.7 | 19.6 | 20.8 | 22.4 | 22.2774 | 23 | 19.8 | 19.8 | 23.5 | 20.92 | 21.35 | 21.37 | 19.35 | 20 | 65.7 | 19.8
| Mean |  | 20.7064 | 21.6533 | 23.15 | 21.8559 | 21.5 | 23.0683 | 18.8683 | 19.4083 | 20.0583 | 20.0083 | 20.3333 | 22.0167 | 22.0862 | 22.5 | 19.2667 | 19.9833 | 23.2 | 21.115 | 21.2517 | 21.265 | 20.6233 | 20.1667 | 63.45 | 19.35
| Median |  | 20.6856 | 21.605 | 22.95 | 21.7623 | 21.5 | 23.355 | 18.735 | 19.325 | 19.95 | 20 | 20.45 | 21.8 | 22.0595 | 23 | 19.25 | 19.95 | 23.25 | 21.19 | 21.2 | 21.455 | 20.66 | 20 | 63.45 | 19.3
| Std Dev. |  | 0.50966 | 0.18283 | 1.3442 | 0.68302 | 0.54772 | 1.0702 | 0.73516 | 0.47267 | 0.90466 | 0.50042 | 0.48339 | 0.52694 | 0.2119 | 0.83666 | 0.41793 | 0.24833 | 0.38987 | 0.7418 | 0.24334 | 0.51968 | 0.94016 | 0.40825 | 1.6072 | 0.29496
| Rel.Std.Dev. |  | 2.461% | 0.8443% | 5.807% | 3.125% | 2.548% | 4.639% | 3.896% | 2.435% | 4.51% | 2.501% | 2.377% | 2.393% | 0.9594% | 3.718% | 2.169% | 1.243% | 1.68% | 3.513% | 1.145% | 2.444% | 4.559% | 2.024% | 2.533% | 1.524%
| PDM3 |  | -1.451% | 3.056% | 10.18% | 4.02% | 2.326% | 9.79% | -10.2% | -7.629% | -4.535% | -4.773% | -3.226% | 4.785% | 5.116% | 7.086% | -8.303% | -4.892% | 10.42% | 0.4939% | 1.144% | 1.208% | -1.846% | -4.02% | 202% | -7.906%
| Z-Score (Absolute) |  | 0.3 | 0.3 | 1.24 | 0.42 | 0.2 | 1.19 | 1.46 | 1.12 | 0.71 | 0.74 | 0.54 | 0.52 | 0.57 | 0.83 | 1.21 | 0.76 | 1.27 | 0.04 | 0.04 | 0.05 | 0.35 | 0.64 | 26.62 | 1.15
| Analytical results for Gd in OREAS 284 (Certified Value 6.19 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.13; 1.12: 1.26; 1.13: 1.29 |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.03: 5.9538; 1.04: 6.03; 1.09: 6.1; 1.10: 6.6572; 1.11: 5.2; 1.12: 6.51; 1.13: 6 |  |  | 1
| 1 | 2 | 1.03: 6.0104; 1.04: 6.04; 1.09: 6.1; 1.10: 6.1901; 1.11: 6.8; 1.12: 6.32; 1.13: 6 |  |  | 37
| 1 | 3 | 1.03: 6.3819; 1.04: 6.14; 1.09: 6.5; 1.10: 6.4494; 1.11: 5.8; 1.12: 6.38; 1.13: 5 |  |  | 16
| 1 | 4 | 1.03: 6.4083; 1.04: 6; 1.09: 6.1; 1.10: 5.9958; 1.11: 5.6; 1.12: 6.55; 1.13: 5 |  |  | 6.1892
| 1 | 5 | 1.03: 6.1092; 1.04: 6.08; 1.09: 5.7; 1.10: 5.6923; 1.11: 5.5; 1.12: 6.43; 1.13: 6 |  |  | 33
| 1 | 6 | 1.03: 6.0219; 1.04: 6.03; 1.09: 6.6; 1.10: 5.8046; 1.11: 5.3; 1.12: 6.39; 1.13: 6
| Mean |  | 1.03: 6.1476; 1.04: 6.0533; 1.09: 6.1833; 1.10: 6.1316; 1.11: 5.7; 1.12: 6.43; 1.13: 5.6667
| Median |  | 1.03: 6.0655; 1.04: 6.035; 1.09: 6.1; 1.10: 6.093; 1.11: 5.55; 1.12: 6.41; 1.13: 6
| Std Dev. |  | 1.03: 0.19823; 1.04: 0.049666; 1.09: 0.32506; 1.10: 0.37391; 1.11: 0.57966; 1.12: 0.086023; 1.13: 0.5164
| Rel.Std.Dev. |  | 1.03: 3.225%; 1.04: 0.8205%; 1.09: 5.257%; 1.10: 6.098%; 1.11: 10.17%; 1.12: 1.338%; 1.13: 9.113%
| PDM3 |  | 1.03: -0.672%; 1.04: -2.195%; 1.09: -0.09415%; 1.10: -0.9306%; 1.11: -7.903%; 1.12: 3.891%; 1.13: -8.442%
| Z-Score (Absolute) |  | 1.03: 0.14; 1.04: 0.67; 1.09: 0.45; 1.10: 0; 1.11: 3.72; 1.12: 2.57; 1.13: 4.01
| Analytical results for Ge in OREAS 284 (Indicative Value 0.27 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | < 0.2 | 1.5 | < 0.1 | 0.33 | 0.08 | 0.11 | 0.15 | 0.12 | 0.11 | 0.38 | 0.57311 | 3.3 | 0.3 | 1.4 | 0.36 | 1.4 | 1.5 | 0.52 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.2 | 1.4 | 0.1 | 0.23 | < 0.05 | 0.12 | 0.15 | 0.12 | 0.1 | 0.5 | 0.62498 | 3.5 | 0.3 | 1.3 | 0.18 | 1.3 | 1.4 | 0.51 | < 0.1 | [7 empty] | 2
| 1 | 3 | < 0.2 | 1.4 | 0.2 | 0.28 | 0.18 | 0.12 | 0.16 | 0.12 | 0.1 | 0.44 | 0.45483 | 3.5 | 0.3 | 1.3 | 0.4 | 1.4 | 1.5 | 0.48 | < 0.1 | [7 empty] | 16
| 1 | 4 | 0.2148 | 1.5 | 0.2 | 0.25 | 0.15 | 0.12 | 0.16 | 0.13 | 0.12 | 0.5 | 0.44484 | 3.7 | 0.4 | 1.3 | 0.27 | 1.2 | 1.5 | 0.54 | < 0.1 | [7 empty] | 0.2682
| 1 | 5 | 0.20604 | 1.3 | < 0.1 | 0.23 | 0.18 | 0.13 | 0.16 | 0.13 | 0.1 | 0.45 | 0.58589 | 3.3 | 0.4 | 1.4 | 0.22 | 1.4 | 1.4 | 0.54 | < 0.1 | [7 empty] | 8
| 1 | 6 | 0.20657 | 1.4 | 0.2 | 0.32 | 0.17 | 0.12 | 0.17 | 0.14 | 0.11 | 0.42 | 0.56122 | 3.2 | 0.3 | 1.3 | 0.38 | 1.4 | 1.5 | 0.55 | < 0.1
| Mean |  | 0.20913 | 1.4167 | 0.175 | 0.27333 | 0.152 | 0.12 | 0.15833 | 0.12667 | 0.10667 | 0.44833 | 0.54081 | 3.4167 | 0.33333 | 1.3333 | 0.30167 | 1.35 | 1.4667 | 0.52333
| Median |  | 0.20657 | 1.4 | 0.2 | 0.265 | 0.17 | 0.12 | 0.16 | 0.125 | 0.105 | 0.445 | 0.56717 | 3.4 | 0.3 | 1.3 | 0.315 | 1.4 | 1.5 | 0.53
| Std Dev. |  | 0.0049104 | 0.075277 | 0.05 | 0.044121 | 0.042071 | 0.0063246 | 0.0075277 | 0.008165 | 0.008165 | 0.046655 | 0.073729 | 0.18348 | 0.05164 | 0.05164 | 0.091305 | 0.083666 | 0.05164 | 0.02582
| Rel.Std.Dev. |  | 2.348% | 5.314% | 28.57% | 16.14% | 27.68% | 5.27% | 4.754% | 6.446% | 7.655% | 10.41% | 13.63% | 5.37% | 15.49% | 3.873% | 30.27% | 6.197% | 3.521% | 4.934%
| PDM3 |  | -22.02% | 428.2% | -34.75% | 1.914% | -43.33% | -55.26% | -40.96% | -52.77% | -60.23% | 67.16% | 101.6% | 1174% | 24.28% | 397.1% | 12.48% | 403.4% | 446.9% | 95.13%
| Z-Score (Absolute) |  | 0.55 | 4.14 | 0.62 | 0.11 | 0.63 | 0.67 | 0.53 | 0.65 | 0.72 | 0.54 | 0.89 | 11.56 | 0.12 | 3.83 | 0 | 3.89 | 4.32 | 0.82 | 0.93
| Analytical results for Hf in OREAS 284 (Certified Value 1.93 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.0181 | 1.93 | 2.4 | 1.67 | 2.05 | 1.9 | 2 | 1.8 | 2 | 2 | 2.1892 | 2.1 | 2.53 | 1.5 | 2 | 1.9 | 2 | 2.04 | 2.02 | 3 | 1.9 | 1.8 |  |  |  |  | 1
| 1 | 2 | 1.9313 | 1.91 | 2.1 | 1.68 | 2.25 | 1.9 | 1.9 | 1.7 | 2.1 | 2 | 2.1029 | 2.1 | 2.52 | 1.6 | 1.8 | 1.7 | 1.96 | 1.91 | 1.98 | 4 | 1.8 | 1.8 |  |  |  |  | 22
| 1 | 3 | 2.0323 | 1.86 | 2.3 | 1.7 | 1.87 | 1.9 | 1.9 | 1.7 | 2.1 | 2 | 1.8024 | 2.1 | 2.48 | 1.7 | 1.9 | 1.7 | 1.84 | 2.11 | 1.88 | 3 | 1.93 | 1.85 |  |  |  |  | 16
| 1 | 4 | 1.9411 | 1.87 | 2.1 | 1.71 | 1.76 | 1.9 | 2 | 1.8 | 2.2 | 2.1 | 2.0067 | 2.2 | 2.43 | 1.6 | 2 | 1.7 | 1.89 | 1.85 | 1.86 | 3 | 1.8 | 1.86 |  |  |  |  | 1.9285
| 1 | 5 | 1.9089 | 1.93 | 2.3 | 1.75 | 1.81 | 2 | 1.9 | 1.7 | 1.9 | 2.1 | 1.8737 | 1.9 | 2.42 | 1.6 | 1.9 | 1.7 | 1.94 | 1.92 | 1.91 | 5 | 1.9 | 1.76 |  |  |  |  | 34
| 1 | 6 | 1.9274 | 1.95 | 2.1 | 1.74 | 1.74 | 1.9 | 2 | 1.8 | 2 | 1.9 | 2.1022 | 2.2 | 2.42 | 1.7 | 2 | 1.7 | 1.9 | 1.85 | 2.1 | 4 | 1.8 | 1.89
| Mean |  | 1.9598 | 1.9083 | 2.2167 | 1.7083 | 1.9133 | 1.9167 | 1.95 | 1.75 | 2.05 | 2.0167 | 2.0129 | 2.1 | 2.4667 | 1.6167 | 1.9333 | 1.7333 | 1.9217 | 1.9467 | 1.9583 | 3.6667 | 1.855 | 1.8267
| Median |  | 1.9362 | 1.92 | 2.2 | 1.705 | 1.84 | 1.9 | 1.95 | 1.75 | 2.05 | 2 | 2.0545 | 2.1 | 2.455 | 1.6 | 1.95 | 1.7 | 1.92 | 1.915 | 1.945 | 3.5 | 1.85 | 1.825
| Std Dev. |  | 0.051873 | 0.036009 | 0.13292 | 0.031885 | 0.19906 | 0.040825 | 0.054772 | 0.054772 | 0.10488 | 0.075277 | 0.14892 | 0.10954 | 0.050465 | 0.075277 | 0.08165 | 0.08165 | 0.056716 | 0.10596 | 0.092177 | 0.8165 | 0.061237 | 0.048028
| Rel.Std.Dev. |  | 2.647% | 1.887% | 5.996% | 1.866% | 10.4% | 2.13% | 2.809% | 3.13% | 5.116% | 3.733% | 7.398% | 5.216% | 2.046% | 4.656% | 4.223% | 4.711% | 2.951% | 5.443% | 4.707% | 22.27% | 3.301% | 2.629%
| PDM3 |  | 1.626% | -1.044% | 14.94% | -11.42% | -0.7852% | -0.6123% | 1.116% | -9.255% | 6.302% | 4.573% | 4.376% | 8.894% | 27.91% | -16.17% | 0.2519% | -10.12% | -0.3531% | 0.9433% | 1.548% | 90.13% | -3.81% | -5.279%
| Z-Score (Absolute) |  | 0.22 | 0.21 | 2.4 | 1.91 | 0.17 | 0.14 | 0.14 | 1.55 | 0.99 | 0.71 | 0.67 | 1.41 | 4.52 | 2.68 | 0 | 1.7 | 0.1 | 0.11 | 0.21 | NA | 0.66 | 0.9
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 284 (Indicative Value 0.13 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 1.03: Lab; 1.04: Lab |  |  |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.12 |  |  |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.03: < 1; 1.04: 0.14 |  |  |  |  | 1
| 1 | 2 | 1.03: < 1; 1.04: 0.13 |  |  |  |  | 5
| 1 | 3 | 1.03: < 1; 1.04: 0.14 |  |  |  |  | 16
| 1 | 4 | 1.03: < 1; 1.04: 0.13 |  |  |  |  | 0.13333
| 1 | 5 | 1.03: < 1; 1.04: 0.13 |  |  |  |  | 7
| 1 | 6 | 1.03: < 1; 1.04: 0.13
| Mean |  | 1.04: 0.13333
| Median |  | 1.04: 0.13
| Std Dev. |  | 1.04: 0.005164
| Rel.Std.Dev. |  | 1.04: 3.873%
| PDM3 |  | 1.04: 2.442e-13%
| Z-Score (Absolute) |  | 1.03: 0.67; 1.04: 0.67
| Analytical results for Ho in OREAS 284 (Certified Value 0.63 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.12: 1.10; 1.13: 1.13; 1.14: 1.26; 1.15: 1.29 |  |  |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS; 1.14: 4A*MS; 1.15: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.03: 0.60284; 1.04: 0.62; 1.09: 0.7; 1.12: 0.4823; 1.13: 0.6; 1.14: 0.61; 1.15: < 1 |  |  |  |  | 1
| 1 | 2 | 1.03: 0.61992; 1.04: 0.61; 1.09: 0.7; 1.12: 0.5174; 1.13: 0.9; 1.14: 0.6; 1.15: < 1 |  |  |  |  | 23
| 1 | 3 | 1.03: 0.62626; 1.04: 0.59; 1.09: 0.7; 1.12: 0.4914; 1.13: 0.6; 1.14: 0.6; 1.15: < 1 |  |  |  |  | 16
| 1 | 4 | 1.03: 0.62668; 1.04: 0.59; 1.09: 0.7; 1.12: 0.536; 1.13: 0.7; 1.14: 0.64; 1.15: < 1 |  |  |  |  | 0.62789
| 1 | 5 | 1.03: 0.61964; 1.04: 0.61; 1.09: 0.6; 1.12: 0.4357; 1.13: 0.6; 1.14: 0.62; 1.15: < 1 |  |  |  |  | 35
| 1 | 6 | 1.03: 0.61148; 1.04: 0.6; 1.09: 0.7; 1.12: 0.479; 1.13: 0.6; 1.14: 0.62; 1.15: < 1
| Mean |  | 1.03: 0.6178; 1.04: 0.60333; 1.09: 0.68333; 1.12: 0.4903; 1.13: 0.66667; 1.14: 0.615
| Median |  | 1.03: 0.61978; 1.04: 0.605; 1.09: 0.7; 1.12: 0.48685; 1.13: 0.6; 1.14: 0.615
| Std Dev. |  | 1.03: 0.0091896; 1.04: 0.012111; 1.09: 0.040825; 1.12: 0.034632; 1.13: 0.12111; 1.14: 0.015166
| Rel.Std.Dev. |  | 1.03: 1.487%; 1.04: 2.007%; 1.09: 5.974%; 1.12: 7.063%; 1.13: 18.17%; 1.14: 2.466%
| PDM3 |  | 1.03: -1.607%; 1.04: -3.912%; 1.09: 8.829%; 1.12: -21.91%; 1.13: 6.175%; 1.14: -2.054%
| Z-Score (Absolute) |  | 1.03: 0.04; 1.04: 0.15; 1.09: 0.89; 1.12: 1.63; 1.13: 0.67; 1.14: 0; 1.15: 1.5
| Analytical results for In in OREAS 284 (Certified Value 0.069 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.07 | < 0.1 | < 2 | 0.092 | 0.06 | 0.071 | 0.078 | 0.062 | 0.071 | 0.076 | 0.070869 | 0.11 | 0.07 | 0.07 | 0.067 | 0.07 | 0.068 | 0.07 | 0.06 | 0.094 | 0.08 | [5 empty] | 1
| 1 | 2 | 0.07 | < 0.1 | < 2 | 0.094 | 0.04 | 0.066 | 0.075 | 0.067 | 0.068 | 0.071 | 0.073296 | 0.09 | 0.07 | 0.07 | 0.065 | 0.06 | 0.072 | 0.07 | 0.08 | 0.093 | 0.09 | [5 empty] | 24
| 1 | 3 | 0.06 | < 0.1 | < 2 | 0.098 | 0.07 | 0.064 | 0.074 | 0.067 | 0.065 | 0.071 | 0.067145 | 0.09 | 0.07 | 0.07 | 0.071 | 0.07 | 0.066 | 0.06 | 0.06 | 0.091 | 0.08 | [5 empty] | 16
| 1 | 4 | 0.07 | < 0.1 | < 2 | 0.095 | 0.08 | 0.069 | 0.074 | 0.071 | 0.07 | 0.077 | 0.065289 | 0.09 | 0.07 | 0.07 | 0.074 | 0.07 | 0.071 | 0.07 | 0.07 | 0.086 | 0.1 | [5 empty] | 0.069018
| 1 | 5 | 0.08 | < 0.1 | < 2 | 0.095 | 0.07 | 0.069 | 0.084 | 0.067 | 0.065 | 0.073 | 0.075823 | 0.09 | 0.07 | 0.07 | 0.069 | 0.07 | 0.073 | 0.07 | 0.07 | 0.096 | 0.09 | [5 empty] | 36
| 1 | 6 | 0.06 | < 0.1 | < 2 | 0.094 | 0.06 | 0.073 | 0.077 | 0.07 | 0.067 | 0.071 | 0.06613 | 0.1 | 0.07 | 0.07 | 0.062 | 0.07 | 0.072 | 0.07 | 0.07 | 0.087 | 0.09
| Mean |  | 0.068333 |  |  | 0.094667 | 0.063333 | 0.068667 | 0.077 | 0.067333 | 0.067667 | 0.073167 | 0.069759 | 0.095 | 0.07 | 0.07 | 0.068 | 0.068333 | 0.070333 | 0.068333 | 0.068333 | 0.091167 | 0.088333
| Median |  | 0.07 |  |  | 0.0945 | 0.065 | 0.069 | 0.076 | 0.067 | 0.0675 | 0.072 | 0.069007 | 0.09 | 0.07 | 0.07 | 0.068 | 0.07 | 0.0715 | 0.07 | 0.07 | 0.092 | 0.09
| Std Dev. |  | 0.0075277 |  |  | 0.0019664 | 0.013663 | 0.003266 | 0.0037947 | 0.0031411 | 0.0025033 | 0.0027142 | 0.0042545 | 0.0083666 | 0 | 0 | 0.0042895 | 0.0040825 | 0.0027325 | 0.0040825 | 0.0075277 | 0.0039707 | 0.0075277
| Rel.Std.Dev. |  | 11.02% |  |  | 2.077% | 21.57% | 4.756% | 4.928% | 4.665% | 3.7% | 3.71% | 6.099% | 8.807% | 0% | 0% | 6.308% | 5.974% | 3.885% | 5.974% | 11.02% | 4.355% | 8.522%
| PDM3 |  | -0.9927% |  |  | 37.16% | -8.237% | -0.5097% | 11.56% | -2.442% | -1.959% | 6.01% | 1.072% | 37.64% | 1.422% | 1.422% | -1.476% | -0.9927% | 1.905% | -0.9927% | -0.9927% | 32.09% | 27.99%
| Z-Score (Absolute) |  | 0.46 | 6.37 | 299.84 | 8.03 | 2.07 | 0.35 | 2.33 | 0.78 | 0.67 | 1.1 | 0 | 8.14 | 0.08 | 0.08 | 0.57 | 0.46 | 0.19 | 0.46 | 0.46 | 6.9 | 5.99
| Analytical results for K in OREAS 284 (Certified Value 2.93 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.7721 | 2.8623 | 2.9445 | 3.14 | 2.9274 | 2.94 | 2.8017 | 2.26 | 2.93 | 3.05 | 2.79 | 2.78 | 2.9 | 2.9399 | 2.88 | 2.97 | 2.67 | 2.79 | 2.86 | 2.88 | 2.9386 | 2.8098 | 2.942 | 3.047 | 1.76 | 2.93 | 1
| 1 | 2 | 2.7735 | 2.859 | 2.9241 | 2.87 | 2.9223 | 2.93 | 2.7966 | 2.48 | 3 | 2.97 | 2.81 | 2.9 | 3.07 | 2.9295 | 2.92 | 3.04 | 2.71 | 2.8 | 2.91 | 3.06 | 2.918 | 2.8998 |  | 3.105 | 1.72 | 2.93 | 0
| 1 | 3 | 2.7565 | 2.938 | 2.9439 | 3.18 | 2.9314 | 2.93 | 2.7984 | 2.45 | 2.89 | 2.96 | 2.8 | 2.9 | 2.99 | 2.9157 | 2.9 | 2.99 | 2.81 | 2.74 | 3.05 | 2.93 | 2.8925 | 2.9038 | 2.778 | 2.989 | 2 | 2.94 | 16
| 1 | 4 | 2.7487 | 2.965 | 2.9641 | 2.86 | 2.9678 | 2.92 | 2.8457 | 2.06 | 2.97 | 2.96 | 2.9 | 2.97 | 3 | 2.9201 | 2.89 | 3.03 | 2.67 | 2.81 | 2.88 | 3.05 | 2.8667 | 2.893 | 2.987 | 2.98 | 1.8 | 3.02 | 2.9261
| 1 | 5 | 2.7461 | 2.8601 | 2.9563 | 3.22 | 2.8548 | 2.94 | 2.7778 | 2.46 | 3.03 | 3.01 | 2.88 | 2.92 | 3.03 | 2.9064 | 2.86 | 3.1 | 2.8 | 2.78 | 2.99 | 3.06 | 2.9823 | 2.9387 | 2.926 | 3.072 | 1.69 | 2.99 | 37
| 1 | 6 | 2.7766 | 2.9157 | 3.0071 | 3.16 | 2.9383 | 2.99 | 2.7896 | 2.07 | 2.96 | 3 | 2.77 | 2.88 | 3.06 | 2.9383 | 2.89 | 3.12 | 2.88 | 2.78 | 2.94 | 3.06 | 2.9168 | 2.93 | 3.049 | 3.171 | 1.78 | 3.01
| Mean |  | 2.7622 | 2.9 | 2.9567 | 3.0717 | 2.9237 | 2.9417 | 2.8016 | 2.2967 | 2.9633 | 2.9917 | 2.825 | 2.8917 | 3.0083 | 2.925 | 2.89 | 3.0417 | 2.7567 | 2.7833 | 2.9383 | 3.0067 | 2.9192 | 2.8958 | 2.9364 | 3.0607 | 1.7917 | 2.97
| Median |  | 2.7643 | 2.889 | 2.9504 | 3.15 | 2.9294 | 2.935 | 2.7975 | 2.355 | 2.965 | 2.985 | 2.805 | 2.9 | 3.015 | 2.9248 | 2.89 | 3.035 | 2.755 | 2.785 | 2.925 | 3.055 | 2.9174 | 2.9018 | 2.942 | 3.0595 | 1.77 | 2.965
| Std Dev. |  | 0.01346 | 0.046061 | 0.02818 | 0.16229 | 0.037363 | 0.024833 | 0.023197 | 0.19623 | 0.049666 | 0.035449 | 0.05244 | 0.062743 | 0.061779 | 0.013228 | 0.02 | 0.059133 | 0.086178 | 0.024221 | 0.071391 | 0.080416 | 0.039584 | 0.045812 | 0.10058 | 0.072246 | 0.10962 | 0.041473
| Rel.Std.Dev. |  | 0.4873% | 1.588% | 0.9531% | 5.283% | 1.278% | 0.8442% | 0.828% | 8.544% | 1.676% | 1.185% | 1.856% | 2.17% | 2.054% | 0.4522% | 0.692% | 1.944% | 3.126% | 0.8702% | 2.43% | 2.675% | 1.356% | 1.582% | 3.425% | 2.36% | 6.118% | 1.396%
| PDM3 |  | -5.599% | -0.8906% | 1.045% | 4.975% | -0.08257% | 0.5326% | -4.253% | -21.51% | 1.273% | 2.241% | -3.455% | -1.176% | 2.811% | -0.03768% | -1.233% | 3.95% | -5.79% | -4.879% | 0.4187% | 2.754% | -0.2369% | -1.033% | 0.3526% | 4.599% | -38.77% | 1.501%
| Z-Score (Absolute) |  | 1.93 | 0.29 | 0.39 | 1.76 | 0.01 | 0.21 | 1.46 | 7.49 | 0.47 | 0.8 | 1.19 | 0.39 | 1 | 0.01 | 0.41 | 1.4 | 2 | 1.68 | 0.17 | 0.98 | 0.06 | 0.34 | 0.14 | 1.62 | 13.52 | 0.55
| Analytical results for La in OREAS 284 (Certified Value 38 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 34.6805 | 39.04 | 37.1 | 38.8504 | 37 | 27.6 | 37.8 | 35.3 | 33 | 38.7 | 33.6 | 38.3 | 44 | 36.5 | 41.8 | 37.9 | 37.03 | 40 | 37.71 | 38 | 27 | 31.3 |  |  |  |  | 1
| 1 | 2 | 34.9458 | 39.63 | 39.3 | 39.4173 | 38 | 28.5 | 37.2 | 38.7 | 38.3 | 35.7 | 36.4 | 37.6 | 44 | 37.2 | 42.5 | 41 | 39.61 | 39 | 38.5 | 39 | 26.5 | 29.1 |  |  |  |  | 25
| 1 | 3 | 35.5638 | 39.36 | 39.7 | 39.8799 | 37 | 28.6 | 40.7 | 35.6 | 38 | 37.7 | 37.2 | 37.3 | 42 | 36.3 | 43.8 | 37.8 | 38.36 | 39 | 38.36 | 38 | 25.9 | 27.7 |  |  |  |  | 16
| 1 | 4 | 34.8363 | 38.68 | 38 | 36.9565 | 37 | 29.4 | 37.7 | 34.8 | 37.8 | 39.6 | 36.3 | 38.5 | 45 | 36.9 | 43.3 | 37.8 | 39.58 | 40 | 38.84 | 38 | 27.3 | 28.4 |  |  |  |  | 37.9609
| 1 | 5 | 35.1082 | 39.41 | 33.9 | 36.2642 | 38 | 29 | 39.6 | 39.4 | 39 | 38.9 | 34.2 | 37.1 | 40 | 35.1 | 43.4 | 40.5 | 40.07 | 40 | 38.73 | 40 | 26.7 | 29.3 |  |  |  |  | 38
| 1 | 6 | 35.0056 | 39.54 | 39.2 | 36.9056 | 38 | 28.6 | 36.7 | 36.3 | 40.6 | 39.2 | 34.6 | 38.1 | 42 | 35 | 44 | 36.5 | 39.81 | 40 | 38.06 | 40 | 28.3 | 28.3
| Mean |  | 35.0233 | 39.2767 | 37.8667 | 38.0457 | 37.5 | 28.6167 | 38.2833 | 36.6833 | 37.7833 | 38.3 | 35.3833 | 37.8167 | 42.8333 | 36.1667 | 43.1333 | 38.5833 | 39.0767 | 39.6667 | 38.3667 | 38.8333 | 26.95 | 29.0167
| Median |  | 34.9757 | 39.385 | 38.6 | 37.9034 | 37.5 | 28.6 | 37.75 | 35.95 | 38.15 | 38.8 | 35.45 | 37.85 | 43 | 36.4 | 43.35 | 37.85 | 39.595 | 40 | 38.43 | 38.5 | 26.85 | 28.75
| Std Dev. |  | 0.30269 | 0.35523 | 2.1695 | 1.5201 | 0.54772 | 0.60139 | 1.5381 | 1.9094 | 2.5538 | 1.4241 | 1.44 | 0.56716 | 1.8348 | 0.92014 | 0.83347 | 1.7634 | 1.1632 | 0.5164 | 0.42406 | 0.98319 | 0.81425 | 1.2592
| Rel.Std.Dev. |  | 0.8642% | 0.9044% | 5.729% | 3.995% | 1.461% | 2.102% | 4.018% | 5.205% | 6.759% | 3.718% | 4.07% | 1.5% | 4.284% | 2.544% | 1.932% | 4.57% | 2.977% | 1.302% | 1.105% | 2.532% | 3.021% | 4.34%
| PDM3 |  | -7.738% | 3.466% | -0.2482% | 0.2233% | -1.214% | -24.62% | 0.8494% | -3.365% | -0.4677% | 0.8933% | -6.79% | -0.3799% | 12.84% | -4.726% | 13.63% | 1.64% | 2.939% | 4.494% | 1.069% | 2.298% | -29.01% | -23.56%
| Z-Score (Absolute) |  | 1.65 | 0.74 | 0.05 | 0.05 | 0.26 | 5.26 | 0.18 | 0.72 | 0.1 | 0.19 | 1.45 | 0.08 | 2.75 | 1.01 | 2.92 | 0.35 | 0.63 | 0.96 | 0.23 | 0.49 | 6.2 | 5.04
| Analytical results for Li in OREAS 284 (Certified Value 53 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 53.6774 | 55.4 | 54 | 87.4666 | 56 | 48 | 48.1 | 54.5 | 58.1 | 51.3 | 51.4 | 52.8 | 54.4684 | 61 | 54 | 47.3 | 49.7 | 58.1 | 51.3 | 54 | 53.8 | 45.47 | 54 | 49.2 | 65 |  | 1
| 1 | 2 | 53.87 | 54.8 | 55 | 82.3583 | 56 | 48 | 52.6 | 55.7 | 56.4 | 50.7 | 53.8 | 55.6 | 54.7112 | 57.5 | 56 | 47.7 | 47.2 | 54.3 | 54.5 | 54 | 53.5 | 50.75 | 54 | 52.2 | 66 |  | 45
| 1 | 3 | 54.472 | 56 | 55 | 82.156 | 55 | 48 | 50.2 | 54.3 | 56.1 | 50.9 | 53.4 | 54.5 | 53.8647 | 57.5 | 57 | 48.7 | 48.7 | 55.6 | 52.2 | 53 | 54.7 | 48.78 | 51 | 50.6 | 64 |  | 16
| 1 | 4 | 53.2179 | 54.6 | 55 | 84.1227 | 58 | 49 | 48 | 55.1 | 56.8 | 52.2 | 55.3 | 53.9 | 53.8913 | 57.7 | 56 | 48.6 | 49.6 | 53.6 | 54.3 | 53 | 54.2 | 46.85 | 52 | 54.5 | 67 |  | 53.169
| 1 | 5 | 52.9078 | 55.2 | 55 | 80.966 | 59 | 48 | 45.7 | 56.4 | 57.2 | 51.8 | 54 | 54.6 | 53.8929 | 54.3 | 57 | 48.5 | 50.2 | 55.3 | 54.8 | 55 | 54.1 | 43.81 | 53 | 54 | 65 |  | 39
| 1 | 6 | 52.8339 | 55.5 | 55 | 83.0028 | 56 | 48 | 45.7 | 55.4 | 57 | 50.5 | 53.6 | 55.7 | 54.6388 | 59.4 | 57 | 49 | 48.4 | 54.1 | 54.6 | 54 | 54.2 |  | 52 | 51.8 | 68
| Mean |  | 53.4965 | 55.25 | 54.8333 | 83.3454 | 56.6667 | 48.1667 | 48.3833 | 55.2333 | 56.9333 | 51.2333 | 53.5833 | 54.5167 | 54.2446 | 57.9 | 56.1667 | 48.3 | 48.9667 | 55.1667 | 53.6167 | 53.8333 | 54.0833 | 47.132 | 52.6667 | 52.05 | 65.8333
| Median |  | 53.4477 | 55.3 | 55 | 82.6805 | 56 | 48 | 48.05 | 55.25 | 56.9 | 51.1 | 53.7 | 54.55 | 54.1807 | 57.6 | 56.5 | 48.55 | 49.15 | 54.8 | 54.4 | 54 | 54.15 | 46.85 | 52.5 | 52 | 65.5
| Std Dev. |  | 0.63042 | 0.50498 | 0.40825 | 2.2696 | 1.5055 | 0.40825 | 2.6724 | 0.77889 | 0.69761 | 0.66232 | 1.2624 | 1.087 | 0.404 | 2.2441 | 1.169 | 0.65422 | 1.093 | 1.6219 | 1.4825 | 0.75277 | 0.40702 | 2.7241 | 1.2111 | 2.0057 | 1.472
| Rel.Std.Dev. |  | 1.178% | 0.914% | 0.7445% | 2.723% | 2.657% | 0.8476% | 5.523% | 1.41% | 1.225% | 1.293% | 2.356% | 1.994% | 0.7448% | 3.876% | 2.081% | 1.354% | 2.232% | 2.94% | 2.765% | 1.398% | 0.7526% | 5.78% | 2.299% | 3.853% | 2.236%
| PDM3 |  | 0.616% | 3.914% | 3.13% | 56.76% | 6.578% | -9.408% | -9.001% | 3.883% | 7.08% | -3.641% | 0.7793% | 2.535% | 2.023% | 8.898% | 5.638% | -9.158% | -7.904% | 3.757% | 0.842% | 1.25% | 1.72% | -11.35% | -0.9447% | -2.105% | 23.82%
| Z-Score (Absolute) |  | 0.19 | 0.39 | 0.25 | 9.7 | 0.86 | 1.96 | 1.89 | 0.38 | 0.95 | 0.95 | 0.17 | 0.14 | 0.05 | 1.27 | 0.69 | 1.92 | 1.7 | 0.36 | 0.15 | 0.08 | 0 | 2.31 | 0.47 | 0.67 | 3.9
| Analytical results for Lu in OREAS 284 (Certified Value 0.16 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.13; 1.12: 1.21; 1.13: 1.25; 1.14: 1.26; 1.15: 1.32 |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*MS; 1.15: 4A*OES/MS |  | 2
2
| 1 | 1 | 1.03: 0.16917; 1.04: 0.17; 1.09: 0.2; 1.10: 0.1588; 1.11: 0.1; 1.12: 0.17; 1.13: 0.17; 1.14: 0.16; 1.15: 0.15 |  | 1
| 1 | 2 | 1.03: 0.17234; 1.04: 0.17; 1.09: 0.2; 1.10: 0.131; 1.11: 0.2; 1.12: 0.18; 1.13: 0.18; 1.14: 0.15; 1.15: 0.14 |  | 26
| 1 | 3 | 1.03: 0.17189; 1.04: 0.16; 1.09: 0.2; 1.10: 0.1274; 1.11: 0.2; 1.12: 0.18; 1.13: 0.18; 1.14: 0.16; 1.15: 0.14 |  | 16
| 1 | 4 | 1.03: 0.1749; 1.04: 0.16; 1.09: 0.2; 1.10: 0.1351; 1.11: 0.2; 1.12: 0.19; 1.13: 0.18; 1.14: 0.17; 1.15: 0.15 |  | 0.16179
| 1 | 5 | 1.03: 0.17059; 1.04: 0.17; 1.09: 0.2; 1.10: 0.1197; 1.11: 0.1; 1.12: 0.19; 1.13: 0.18; 1.14: 0.16; 1.15: 0.14 |  | 40
| 1 | 6 | 1.03: 0.16692; 1.04: 0.16; 1.09: 0.2; 1.10: 0.1278; 1.11: 0.1; 1.12: 0.19; 1.13: 0.18; 1.14: 0.16; 1.15: 0.16
| Mean |  | 1.03: 0.17097; 1.04: 0.165; 1.09: 0.2; 1.10: 0.1333; 1.11: 0.15; 1.12: 0.18333; 1.13: 0.17833; 1.14: 0.16; 1.15: 0.14667
| Median |  | 1.03: 0.17124; 1.04: 0.165; 1.09: 0.2; 1.10: 0.1294; 1.11: 0.15; 1.12: 0.185; 1.13: 0.18; 1.14: 0.16; 1.15: 0.145
| Std Dev. |  | 1.03: 0.0027546; 1.04: 0.0054772; 1.09: 3.0405e-17; 1.10: 0.013482; 1.11: 0.054772; 1.12: 0.008165; 1.13: 0.0040825; 1.14: 0.0063246; 1.15: 0.008165
| Rel.Std.Dev. |  | 1.03: 1.611%; 1.04: 3.32%; 1.09: 1.52e-14%; 1.10: 10.11%; 1.11: 36.51%; 1.12: 4.454%; 1.13: 2.289%; 1.14: 3.953%; 1.15: 5.567%
| PDM3 |  | 1.03: 5.675%; 1.04: 1.987%; 1.09: 23.62%; 1.10: -17.61%; 1.11: -7.285%; 1.12: 13.32%; 1.13: 10.23%; 1.14: -1.104%; 1.15: -9.345%
| Z-Score (Absolute) |  | 1.03: 0.3; 1.04: 0; 1.09: NA; 1.10: 1.6; 1.11: NA; 1.12: 0.93; 1.13: 0.67; 1.14: 0.25; 1.15: 0.93
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 284 (Certified Value 0.852 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.89599 | 0.82709 | 0.8286 | 0.89 | 0.8053 | 0.85 | 0.8527 | 0.87 | 0.8 | 0.87 | 0.79 | 0.77 | 0.81 | 0.86022 | 0.82 | 0.76 | 0.855 | 0.87 | 0.83 | 0.819 | 0.8611 | 0.8481 | 0.881 | 0.911 | 0.85 | 0.87 | 1
| 1 | 2 | 0.90587 | 0.83631 | 0.8283 | 0.85 | 0.7957 | 0.83 | 0.8558 | 0.84 | 0.83 | 0.86 | 0.78 | 0.81 | 0.86 | 0.8644 | 0.84 | 0.77 | 0.886 | 0.87 | 0.83 | 0.864 | 0.8557 | 0.8554 | 0.967 | 0.935 | 0.87 | 0.86 | 0
| 1 | 3 | 0.90467 | 0.8317 | 0.8555 | 0.89 | 0.7964 | 0.83 | 0.8503 | 0.88 | 0.8 | 0.85 | 0.78 | 0.82 | 0.85 | 0.85681 | 0.82 | 0.78 | 0.862 | 0.85 | 0.83 | 0.829 | 0.8524 | 0.8643 |  | 0.905 | 0.851 | 0.85 | 16
| 1 | 4 | 0.89748 | 0.84893 | 0.838 | 0.89 | 0.8065 | 0.84 | 0.8697 | 0.83 | 0.81 | 0.85 | 0.81 | 0.83 | 0.85 | 0.85763 | 0.84 | 0.76 | 0.883 | 0.89 | 0.82 | 0.865 | 0.8427 | 0.873 | 0.905 | 0.911 | 0.9 | 0.88 | 0.85244
| 1 | 5 | 0.89661 | 0.83198 | 0.8402 | 0.86 | 0.7799 | 0.84 | 0.8438 | 0.86 | 0.83 | 0.87 | 0.8 | 0.81 | 0.85 | 0.86635 | 0.81 | 0.79 | 0.864 | 0.86 | 0.84 | 0.869 | 0.8781 | 0.8729 | 0.911 | 0.935 | 0.89 | 0.86 | 41
| 1 | 6 | 0.90326 | 0.83861 | 0.852 | 0.9 | 0.7983 | 0.84 | 0.853 | 0.83 | 0.81 | 0.86 | 0.78 | 0.8 | 0.86 | 0.85584 | 0.82 | 0.78 | 0.87 | 0.86 | 0.83 | 0.868 | 0.8543 | 0.8753 | 0.864 | 0.947 | 0.92 | 0.85
| Mean |  | 0.90065 | 0.83577 | 0.84043 | 0.88 | 0.79702 | 0.83833 | 0.85422 | 0.85167 | 0.81333 | 0.86 | 0.79 | 0.80667 | 0.84667 | 0.86021 | 0.825 | 0.77333 | 0.87 | 0.86667 | 0.83 | 0.85233 | 0.85738 | 0.86483 | 0.9056 | 0.924 | 0.88017 | 0.86167
| Median |  | 0.90037 | 0.83415 | 0.8391 | 0.89 | 0.79735 | 0.84 | 0.85285 | 0.85 | 0.81 | 0.86 | 0.785 | 0.81 | 0.85 | 0.85893 | 0.82 | 0.775 | 0.867 | 0.865 | 0.83 | 0.8645 | 0.855 | 0.8686 | 0.905 | 0.923 | 0.88 | 0.86
| Std Dev. |  | 0.0044354 | 0.0075874 | 0.011437 | 0.02 | 0.0095424 | 0.0075277 | 0.0086008 | 0.02137 | 0.013663 | 0.0089443 | 0.012649 | 0.020656 | 0.018619 | 0.0043038 | 0.012247 | 0.012111 | 0.012247 | 0.013663 | 0.0063246 | 0.02225 | 0.011796 | 0.011051 | 0.039164 | 0.017146 | 0.028074 | 0.01169
| Rel.Std.Dev. |  | 0.4925% | 0.9078% | 1.361% | 2.273% | 1.197% | 0.8979% | 1.007% | 2.509% | 1.68% | 1.04% | 1.601% | 2.561% | 2.199% | 0.5003% | 1.485% | 1.566% | 1.408% | 1.576% | 0.762% | 2.61% | 1.376% | 1.278% | 4.325% | 1.856% | 3.19% | 1.357%
| PDM3 |  | 5.656% | -1.955% | -1.408% | 3.233% | -6.502% | -1.655% | 0.2086% | -0.0905% | -4.587% | 0.8871% | -7.325% | -5.369% | -0.6771% | 0.9117% | -3.219% | -9.28% | 2.06% | 1.669% | -2.632% | -0.01229% | 0.5801% | 1.454% | 6.236% | 8.395% | 3.253% | 1.083%
| Z-Score (Absolute) |  | 1.87 | 0.69 | 0.51 | 1.05 | 2.22 | 0.59 | 0.04 | 0.06 | 1.57 | 0.26 | 2.49 | 1.84 | 0.26 | 0.27 | 1.11 | 3.15 | 0.66 | 0.53 | 0.92 | 0.04 | 0.16 | 0.46 | 2.06 | 2.77 | 1.06 | 0.33
| Analytical results for Mn in OREAS 284 (Certified Value 0.038 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.036743 | 0.0373 | 0.0389 | 0.0511 | 0.0419 | 0.0349 | 0.038 | 0.0364 | 0.0395 | 0.0357 | 0.0354 | 0.0363 | 0.039181 | 0.039 | 0.0349 | 0.0353 | 0.0376 | 0.0406 | 0.035 | 0.0404 | 0.0374 | 0.039 | 0.0471 | 0.04 |  |  | 1
| 1 | 2 | 0.036788 | 0.037 | 0.0387 | 0.0523 | 0.0411 | 0.035 | 0.0401 | 0.0369 | 0.0385 | 0.035 | 0.0369 | 0.0393 | 0.039307 | 0.0384 | 0.0356 | 0.0345 | 0.0374 | 0.0357 | 0.0369 | 0.0386 | 0.0385 | 0.039 | 0.0475 | 0.04 |  |  | 27
| 1 | 3 | 0.037225 | 0.0382 | 0.0401 | 0.0525 | 0.0414 | 0.035 | 0.0388 | 0.0359 | 0.0386 | 0.0349 | 0.0376 | 0.0382 | 0.037487 | 0.0388 | 0.0364 | 0.0355 | 0.0365 | 0.0376 | 0.0355 | 0.0384 | 0.0386 | 0.039 | 0.0455 | 0.04 |  |  | 16
| 1 | 4 | 0.037998 | 0.0377 | 0.0398 | 0.0541 | 0.0418 | 0.0356 | 0.0367 | 0.0363 | 0.0389 | 0.0364 | 0.0382 | 0.0384 | 0.038008 | 0.0392 | 0.0363 | 0.034 | 0.0379 | 0.0367 | 0.0371 | 0.0383 | 0.039 | 0.039 | 0.0479 | 0.04 |  |  | 0.037781
| 1 | 5 | 0.036962 | 0.0375 | 0.0406 | 0.053 | 0.0414 | 0.0348 | 0.0381 | 0.0373 | 0.0391 | 0.0361 | 0.0372 | 0.0385 | 0.037536 | 0.0354 | 0.037 | 0.0354 | 0.037 | 0.0377 | 0.0373 | 0.0399 | 0.0388 | 0.039 | 0.0469 | 0.04 |  |  | 42
| 1 | 6 | 0.037226 | 0.0374 | 0.0383 | 0.0541 | 0.0417 | 0.0348 | 0.0367 | 0.0365 | 0.0392 | 0.0353 | 0.0368 | 0.0392 | 0.039094 | 0.0381 | 0.0373 | 0.0358 | 0.0369 | 0.0373 | 0.0371 | 0.0386 | 0.0387 | 0.039 | 0.0486 | 0.04
| Mean |  | 0.037157 | 0.037517 | 0.0394 | 0.05285 | 0.04155 | 0.035017 | 0.038067 | 0.03655 | 0.038967 | 0.035567 | 0.037017 | 0.038317 | 0.038435 | 0.03815 | 0.03625 | 0.035083 | 0.037217 | 0.0376 | 0.036483 | 0.039033 | 0.0385 | 0.039 | 0.04725 | 0.04
| Median |  | 0.037094 | 0.03745 | 0.03935 | 0.05275 | 0.04155 | 0.03495 | 0.03805 | 0.03645 | 0.039 | 0.0355 | 0.03705 | 0.03845 | 0.038551 | 0.0386 | 0.03635 | 0.03535 | 0.0372 | 0.03745 | 0.037 | 0.0386 | 0.03865 | 0.039 | 0.0473 | 0.04
| Std Dev. |  | 0.00046117 | 0.00040702 | 0.00089889 | 0.001152 | 0.00030166 | 0.00029944 | 0.0012972 | 0.00048888 | 0.00037771 | 0.00060553 | 0.00094322 | 0.0010834 | 0.0008533 | 0.0014053 | 0.00088713 | 0.00068532 | 0.00051153 | 0.0016444 | 0.00097656 | 0.00088694 | 0.00056569 | 0 | 0.0010502 | 0
| Rel.Std.Dev. |  | 1.241% | 1.085% | 2.281% | 2.18% | 0.726% | 0.8551% | 3.408% | 1.338% | 0.9693% | 1.703% | 2.548% | 2.827% | 2.22% | 3.684% | 2.447% | 1.953% | 1.374% | 4.373% | 2.677% | 2.272% | 1.469% | 0% | 2.223% | 0%
| PDM3 |  | -1.65% | -0.6984% | 4.287% | 39.89% | 9.977% | -7.316% | 0.7573% | -3.257% | 3.14% | -5.86% | -2.022% | 1.419% | 1.733% | 0.9779% | -4.051% | -7.139% | -1.492% | -0.4779% | -3.434% | 3.316% | 1.904% | 3.228% | 25.06% | 5.875%
| Z-Score (Absolute) |  | 0.63 | 0.39 | 0.85 | 9.73 | 2.27 | 2.04 | 0.03 | 1.03 | 0.57 | 1.68 | 0.72 | 0.14 | 0.22 | 0.03 | 1.23 | 2 | 0.59 | 0.34 | 1.07 | 0.61 | 0.26 | 0.41 | 6.03 | 1.25
| Analytical results for Mo in OREAS 284 (Certified Value 2.58 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.5514 | 2.5 | 3 | 3.0808 | 3 | 2.8 | 2.16 | 2.41 | 2.45 | 2.53 | 2.45 | 2.63 | 2.6255 | 2.1 | 2.68 | 2.4 | 2.59 | 2.6 | 2.3 | 2.4 | 2.8 | 3 | 3.11 | 2.46 |  |  | 1
| 1 | 2 | 2.4473 | 2.5 | 3 | 2.886 | 3 | 3 | 2.34 | 2.35 | 2.66 | 2.5 | 3.01 | 3.51 | 2.6655 | 2 | 2.66 | 2.5 | 2.7 | 2.3 | 2.44 | 2.6 | 2.5 | 4 | 3.05 | 2.76 |  |  | 28
| 1 | 3 | 2.5827 | 2.8 | 2 | 1.9515 | 3 | 2.9 | 2.37 | 2.13 | 2.54 | 2.53 | 2.49 | 2.66 | 2.4356 | 2 | 2.89 | 2.4 | 2.66 | 2.4 | 2.42 | 2.8 | 2.6 | 4 | 3.42 | 2.36 |  |  | 16
| 1 | 4 | 2.5396 | 2.7 | < 1 | 2.0236 | 3 | 3 | 2.47 | 2.42 | 2.6 | 2.7 | 2.4 | 2.67 | 2.435 | 2.2 | 2.77 | 2.4 | 2.71 | 2.4 | 2.56 | 2.4 | 2.6 | 1 | 3.16 | 2.58 |  |  | 2.5769
| 1 | 5 | 2.6892 | 2.7 | 2 | 2.502 | 3 | 2.6 | 2.08 | 2.6 | 2.62 | 2.81 | 2.43 | 2.63 | 2.6326 | 1.9 | 2.64 | 2.9 | 2.9 | 2.5 | 2.47 | 2.8 | 2.5 | 5 | 3.09 | 2.61 |  |  | 43
| 1 | 6 | 2.5151 | 2.7 | 1 | 3.3345 | 3 | 2.8 | 2.52 | 2.54 | 2.98 | 2.63 | 2.66 | 2.52 | 2.7151 | 2.1 | 2.76 | 2.9 | 2.62 | 2.6 | 2.77 | 2.6 | 3 | 6 | 3.01 | 2.61
| Mean |  | 2.5542 | 2.65 | 2.2 | 2.6297 | 3 | 2.85 | 2.3233 | 2.4083 | 2.6417 | 2.6167 | 2.5733 | 2.77 | 2.5849 | 2.05 | 2.7333 | 2.5833 | 2.6967 | 2.4667 | 2.4933 | 2.6 | 2.6667 | 3.8333 | 3.14 | 2.5633
| Median |  | 2.5455 | 2.7 | 2 | 2.694 | 3 | 2.85 | 2.355 | 2.415 | 2.61 | 2.58 | 2.47 | 2.645 | 2.629 | 2.05 | 2.72 | 2.45 | 2.68 | 2.45 | 2.455 | 2.6 | 2.6 | 4 | 3.1 | 2.595
| Std Dev. |  | 0.080286 | 0.12247 | 0.83666 | 0.56737 | 0 | 0.15166 | 0.17236 | 0.16437 | 0.18115 | 0.12094 | 0.23278 | 0.36644 | 0.12011 | 0.10488 | 0.09331 | 0.24833 | 0.10967 | 0.12111 | 0.15946 | 0.17889 | 0.19664 | 1.7224 | 0.14642 | 0.13808
| Rel.Std.Dev. |  | 3.143% | 4.622% | 38.03% | 21.58% | 0% | 5.321% | 7.418% | 6.825% | 6.858% | 4.622% | 9.046% | 13.23% | 4.647% | 5.116% | 3.414% | 9.613% | 4.067% | 4.91% | 6.395% | 6.88% | 7.374% | 44.93% | 4.663% | 5.387%
| PDM3 |  | -0.8814% | 2.836% | -14.63% | 2.049% | 16.42% | 10.6% | -9.841% | -6.542% | 2.512% | 1.542% | -0.1392% | 7.493% | 0.3079% | -20.45% | 6.07% | 0.2488% | 4.647% | -4.279% | -3.244% | 0.8956% | 3.483% | 48.76% | 21.85% | -0.5273%
| Z-Score (Absolute) |  | 0.46 | 0.51 | NA | 0.3 | NA | 2.53 | 2.8 | 1.94 | 0.42 | 0.17 | 0.27 | 1.72 | 0.15 | 5.56 | 1.35 | 0.17 | 0.98 | 1.35 | 1.08 | 0 | 0.67 | NA | 5.46 | 0.37
Results from laboratories 1.09, 1.11 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 284 (Certified Value 2.04 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.9598 | 2.0038 | 2.0116 | 2.12 | 1.9838 | 2.01 | 2.1699 | 2.128 | 2.02 | 2.12 | 1.97 | 1.87 | 2 | 2.0309 | 1.92 | 2.08 | 1.64 | 1.97 | 1.997 | 1.953 | 2.0693 | 2.0365 | 1.569 | 2.077 | 1.92 | 2.29 | 1
| 1 | 2 | 1.9767 | 2.0096 | 1.9919 | 2.14 | 1.9815 | 1.99 | 2.171 | 2.026 | 2.06 | 2.08 | 1.96 | 1.96 | 2.12 | 2.0555 | 1.93 | 2.13 | 1.66 | 1.97 | 2.029 | 2.069 | 2.0484 | 2.062 | 1.704 | 2.114 | 1.97 | 2.3 | 0
| 1 | 3 | 1.9978 | 2.0375 | 2.0457 | 2.23 | 1.9775 | 1.99 | 2.1568 | 2.156 | 2.01 | 2.05 | 1.96 | 1.98 | 2.08 | 2.0217 | 1.93 | 2.18 | 1.72 | 1.92 | 2.111 | 1.978 | 2.0356 | 2.0759 |  | 2.04 | 1.95 | 2.27 | 16
| 1 | 4 | 1.95 | 2.0359 | 2.0091 | 2.19 | 2.0112 | 2.03 | 2.2093 | 2.051 | 2.03 | 2.07 | 2 | 2.01 | 2.08 | 2.0218 | 1.93 | 2.12 | 1.62 | 2 | 2.063 | 2.066 | 2.0176 | 2.0831 | 1.604 | 2.025 | 2.02 | 2.34 | 2.0414
| 1 | 5 | 1.9793 | 1.9802 | 2.0108 | 2.15 | 1.9608 | 2.01 | 2.1488 | 2.165 | 2.07 | 2.09 | 2 | 1.97 | 2.1 | 2.0531 | 1.83 | 2.17 | 1.69 | 1.95 | 2.09 | 2.064 | 2.0962 | 2.1021 | 1.676 | 2.099 | 2 | 2.29 | 44
| 1 | 6 | 1.9872 | 1.9838 | 2.0418 | 2.13 | 2.0053 | 2.02 | 2.1601 | 2.058 | 2.04 | 2.09 | 1.95 | 1.96 | 2.12 | 2.0311 | 1.93 | 2.14 | 1.75 | 1.96 | 2.075 | 2.064 | 2.0422 | 2.0966 | 1.799 | 2.174 | 2.07 | 2.32
| Mean |  | 1.9751 | 2.0085 | 2.0185 | 2.16 | 1.9867 | 2.0083 | 2.1693 | 2.0973 | 2.0383 | 2.0833 | 1.9733 | 1.9583 | 2.0833 | 2.0357 | 1.9117 | 2.1367 | 1.68 | 1.9617 | 2.0608 | 2.0323 | 2.0516 | 2.076 | 1.6704 | 2.0882 | 1.9883 | 2.3017
| Median |  | 1.978 | 2.0067 | 2.0112 | 2.145 | 1.9827 | 2.01 | 2.165 | 2.093 | 2.035 | 2.085 | 1.965 | 1.965 | 2.09 | 2.031 | 1.93 | 2.135 | 1.675 | 1.965 | 2.069 | 2.064 | 2.0453 | 2.0795 | 1.676 | 2.088 | 1.985 | 2.295
| Std Dev. |  | 0.017591 | 0.024604 | 0.020907 | 0.041952 | 0.018639 | 0.016021 | 0.021281 | 0.059571 | 0.023166 | 0.023381 | 0.021602 | 0.047081 | 0.044572 | 0.015045 | 0.040208 | 0.036148 | 0.049396 | 0.026394 | 0.041619 | 0.052401 | 0.027615 | 0.024134 | 0.089985 | 0.054013 | 0.053448 | 0.024833
| Rel.Std.Dev. |  | 0.8906% | 1.225% | 1.036% | 1.942% | 0.9382% | 0.7977% | 0.981% | 2.84% | 1.137% | 1.122% | 1.095% | 2.404% | 2.139% | 0.7391% | 2.103% | 1.692% | 2.94% | 1.346% | 2.02% | 2.578% | 1.346% | 1.163% | 5.387% | 2.587% | 2.688% | 1.079%
| PDM3 |  | -3.249% | -1.616% | -1.125% | 5.807% | -2.683% | -1.622% | 6.264% | 2.738% | -0.1526% | 2.052% | -3.337% | -4.071% | 2.052% | -0.2819% | -6.357% | 4.664% | -17.71% | -3.908% | 0.9496% | -0.4465% | 0.4948% | 1.694% | -18.18% | 2.289% | -2.602% | 12.75%
| Z-Score (Absolute) |  | 0.77 | 0.33 | 0.2 | 1.64 | 0.62 | 0.33 | 1.76 | 0.82 | 0.06 | 0.64 | 0.79 | 0.99 | 0.64 | 0.02 | 1.59 | 1.34 | 4.61 | 0.94 | 0.35 | 0.02 | 0.23 | 0.55 | 4.73 | 0.71 | 0.59 | 3.48
| Analytical results for Nb in OREAS 284 (Certified Value 12.9 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 13.699 | 13.41 | 7.3 | 12.608 | 16 | 16.39 | 11.86 | 13.5 | 13.7 | 12.6 | 11.7 | 14 | 13.2771 | 13.9 | 13.2 | 11.8 | 13.3 | 13.6 | 12.1 | 12.76 | 12.82 | 10 | 12.8 | 12.3 |  |  | 1
| 1 | 2 | 13.2709 | 13.31 | 3.3 | 11.2019 | 15 | 14.95 | 14.11 | 13.5 | 13.3 | 11.8 | 12.4 | 14.2 | 13.2687 | 13.7 | 13.3 | 11.8 | 13.1 | 11.9 | 13.3 | 12.61 | 13 | 10 | 13.2 | 12.5 |  |  | 29
| 1 | 3 | 13.2075 | 13.45 | 2.9 | 10.5635 | 14 | 15.87 | 12.33 | 13.7 | 13.2 | 12.5 | 12.2 | 13.9 | 12.9814 | 13.8 | 13.7 | 11.2 | 13.4 | 12.9 | 12.7 | 12.33 | 13.09 | 9 | 12.8 | 12.4 |  |  | 16
| 1 | 4 | 13.2506 | 13.19 | 2.1 | 5.8975 | 14 | 16.77 | 12.31 | 13.9 | 13.1 | 12.6 | 12.3 | 14.8 | 13.156 | 13.3 | 13.3 | 11.9 | 13.4 | 11.8 | 13.2 | 12.42 | 13.18 | 11 | 13.4 | 12.6 |  |  | 12.8932
| 1 | 5 | 13.3971 | 13.25 | 12.7 | 10.0301 | 14 | 14.18 | 11.6 | 14 | 13.5 | 12.4 | 11.8 | 13.8 | 13.0918 | 13.1 | 13.4 | 11.6 | 13.4 | 12.2 | 13.2 | 12.91 | 13.2 | 11 | 13.3 | 12.5 |  |  | 45
| 1 | 6 | 12.9584 | 13.39 | 1.2 | 12.1406 | 14 | 16 | 11.35 | 14 | 13.5 | 12.1 | 12.4 | 14.4 | 13.0936 | 13.7 | 13.3 | 11.3 | 13.5 | 12 | 13.4 | 12.69 | 13.15 | 11 | 13.8 | 12.5
| Mean |  | 13.2973 | 13.3333 | 4.9167 | 10.4069 | 14.5 | 15.6933 | 12.26 | 13.7667 | 13.3833 | 12.3333 | 12.1333 | 14.1833 | 13.1448 | 13.5833 | 13.3667 | 11.6 | 13.35 | 12.4 | 12.9833 | 12.62 | 13.0733 | 10.3333 | 13.2167 | 12.4667
| Median |  | 13.2608 | 13.35 | 3.1 | 10.8827 | 14 | 15.935 | 12.085 | 13.8 | 13.4 | 12.45 | 12.25 | 14.1 | 13.1248 | 13.7 | 13.3 | 11.7 | 13.4 | 12.1 | 13.2 | 12.65 | 13.12 | 10.5 | 13.25 | 12.5
| Std Dev. |  | 0.24376 | 0.10073 | 4.3517 | 2.4076 | 0.83666 | 0.96047 | 0.98525 | 0.23381 | 0.22286 | 0.32042 | 0.30768 | 0.37103 | 0.1141 | 0.31252 | 0.17512 | 0.28983 | 0.13784 | 0.70711 | 0.49565 | 0.21578 | 0.14362 | 0.8165 | 0.38166 | 0.10328
| Rel.Std.Dev. |  | 1.833% | 0.7555% | 88.51% | 23.14% | 5.77% | 6.12% | 8.036% | 1.698% | 1.665% | 2.598% | 2.536% | 2.616% | 0.868% | 2.301% | 1.31% | 2.499% | 1.033% | 5.702% | 3.818% | 1.71% | 1.099% | 7.902% | 2.888% | 0.8284%
| PDM3 |  | 3.134% | 3.414% | -61.87% | -19.28% | 12.46% | 21.72% | -4.911% | 6.775% | 3.802% | -4.342% | -5.893% | 10.01% | 1.951% | 5.353% | 3.672% | -10.03% | 3.543% | -3.825% | 0.6991% | -2.119% | 1.397% | -19.85% | 2.509% | -3.308%
| Z-Score (Absolute) |  | 0.22 | 0.27 | 9.77 | 3.22 | NA | 3.08 | 1.01 | 0.78 | 0.33 | 0.92 | 1.16 | 1.28 | 0.04 | 0.57 | 0.31 | 1.8 | 0.29 | 0.85 | 0.15 | 0.58 | 0.04 | NA | 0.13 | 0.77
Results from laboratories 1.11 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 284 (Certified Value 35.6 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.13; 1.12: 1.26; 1.13: 1.29 |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*OES/MS |  |  | 1
2
| 1 | 1 | 1.03: 34.1049; 1.04: 36.52; 1.09: 34.6; 1.10: 35.7987; 1.11: 32.1; 1.12: 37.46; 1.13: 44 |  |  | 1
| 1 | 2 | 1.03: 32.8179; 1.04: 37.46; 1.09: 34.5; 1.10: 34.7548; 1.11: 41.1; 1.12: 37.1; 1.13: 45 |  |  | 6
| 1 | 3 | 1.03: 35.151; 1.04: 36.74; 1.09: 35.6; 1.10: 35.3961; 1.11: 36.2; 1.12: 37.05; 1.13: 44 |  |  | 16
| 1 | 4 | 1.03: 35.8016; 1.04: 36.66; 1.09: 34.6; 1.10: 32.9024; 1.11: 36.4; 1.12: 38.7; 1.13: 44 |  |  | 35.6018
| 1 | 5 | 1.03: 34.6949; 1.04: 36.74; 1.09: 32; 1.10: 31.6066; 1.11: 34; 1.12: 38.35; 1.13: 45 |  |  | 46
| 1 | 6 | 1.03: 34.2791; 1.04: 37.2; 1.09: 35.8; 1.10: 32.9077; 1.11: 32.7; 1.12: 38.85; 1.13: 45
| Mean |  | 1.03: 34.4749; 1.04: 36.8867; 1.09: 34.5167; 1.10: 33.8944; 1.11: 35.4167; 1.12: 37.9183; 1.13: 44.5
| Median |  | 1.03: 34.487; 1.04: 36.74; 1.09: 34.6; 1.10: 33.8312; 1.11: 35.1; 1.12: 37.905; 1.13: 44.5
| Std Dev. |  | 1.03: 1.0185; 1.04: 0.36214; 1.09: 1.3541; 1.10: 1.6621; 1.11: 3.2945; 1.12: 0.81229; 1.13: 0.54772
| Rel.Std.Dev. |  | 1.03: 2.954%; 1.04: 0.9818%; 1.09: 3.923%; 1.10: 4.904%; 1.11: 9.302%; 1.12: 2.142%; 1.13: 1.231%
| PDM3 |  | 1.03: -3.165%; 1.04: 3.609%; 1.09: -3.048%; 1.10: -4.796%; 1.11: -0.5201%; 1.12: 6.507%; 1.13: 24.99%
| Z-Score (Absolute) |  | 1.03: 0.43; 1.04: 0.67; 1.09: 0.41; 1.10: 0.7; 1.11: 0; 1.12: 1.15; 1.13: 4.17
| Analytical results for Ni in OREAS 284 (Certified Value 18.9 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 17.55 | 18.7462 | 19.8 | 20 | 29.6908 | 21 | 18.6 | 18.2 | 19.2 | 18.2 | 18.5 | 18.2 | 18.2 | 19.4081 | 16.2 | 17.6 | 19 | 18.8 | 20.4 | 17 | 20.7 | 19.5 | 18.38 | 27 | 18.7 | 19.5 | 1
| 1 | 2 | 17.61 | 18.6234 | 19.5 | 20 | 24.1627 | 21 | 18.7 | 17.5 | 20.5 | 17.8 | 17.4 | 19.4 | 19.6 | 19.6607 | 15.6 | 17.9 | 19 | 18.9 | 18.9 | 18 | 20.3 | 19.8 |  | 24 | 18 | 19.9 | 16
| 1 | 3 | 17.95 | 19.5065 | 19.8 | 20 | 24.5641 | 21 | 18.5 | 18.3 | 18.8 | 17.6 | 18.2 | 19.4 | 18.2 | 19.5619 | 15.8 | 17.9 | 20 | 18.4 | 19.1 | 17 | 20.3 | 19.6 | 18.4 | 24 | 18.5 | 19.2 | 16
| 1 | 4 | 18.1 | 19.2054 | 19.7 | 20 | 24.6225 | 22 | 18.8 | 17.2 | 19.2 | 19.7 | 19.1 | 19.4 | 19.2 | 19.2276 | 15.3 | 17.9 | 19 | 19.5 | 19 | 18 | 20.5 | 19.6 | 17.91 | 23 | 16.7 | 19.6 | 18.8507
| 1 | 5 | 17.53 | 19.25 | 19.7 | 20 | 21.5793 | 21 | 18.5 | 17.8 | 19.9 | 18.2 | 19.4 | 18.4 | 18.6 | 19.128 | 14.9 | 18.3 | 20 | 18.6 | 19.7 | 18 | 20.7 | 19.9 | 17.48 | 26 | 18.1 | 19.6 | 47
| 1 | 6 | 18.32 | 18.5959 | 19.7 | 22 | 23.7859 | 21 | 18.5 | 17 | 19.5 | 19.6 | 18.6 | 19.6 | 19 | 19.5081 | 16 | 18.9 | 19 | 18.4 | 18.9 | 18 | 20.2 | 19.6 | 16.88 | 23 | 19 | 19.9
| Mean |  | 17.8433 | 18.9879 | 19.7 | 20.3333 | 24.7342 | 21.1667 | 18.6 | 17.6667 | 19.5167 | 18.5167 | 18.5333 | 19.0667 | 18.8 | 19.4157 | 15.6333 | 18.0833 | 19.3333 | 18.7667 | 19.3333 | 17.6667 | 20.45 | 19.6667 | 17.81 | 24.5 | 18.1667 | 19.6167
| Median |  | 17.78 | 18.9758 | 19.7 | 20 | 24.3634 | 21 | 18.55 | 17.65 | 19.35 | 18.2 | 18.55 | 19.4 | 18.8 | 19.4581 | 15.7 | 17.9 | 19 | 18.7 | 19.05 | 18 | 20.4 | 19.6 | 17.91 | 24 | 18.3 | 19.6
| Std Dev. |  | 0.32959 | 0.38205 | 0.10954 | 0.8165 | 2.6754 | 0.40825 | 0.12649 | 0.52789 | 0.6047 | 0.90866 | 0.70333 | 0.60222 | 0.56569 | 0.20403 | 0.4761 | 0.45789 | 0.5164 | 0.41312 | 0.60222 | 0.5164 | 0.21679 | 0.15055 | 0.64358 | 1.6432 | 0.80911 | 0.26394
| Rel.Std.Dev. |  | 1.847% | 2.012% | 0.5561% | 4.016% | 10.82% | 1.929% | 0.6801% | 2.988% | 3.098% | 4.907% | 3.795% | 3.158% | 3.009% | 1.051% | 3.045% | 2.532% | 2.671% | 2.201% | 3.115% | 2.923% | 1.06% | 0.7655% | 3.614% | 6.707% | 4.454% | 1.346%
| PDM3 |  | -5.344% | 0.7277% | 4.505% | 7.865% | 31.21% | 12.29% | -1.33% | -6.281% | 3.533% | -1.772% | -1.684% | 1.146% | -0.269% | 2.997% | -17.07% | -4.071% | 2.56% | -0.4459% | 2.56% | -6.281% | 8.484% | 4.328% | -5.521% | 29.97% | -3.629% | 4.063%
| Z-Score (Absolute) |  | 0.9 | 0.18 | 0.85 | NA | 5.58 | NA | 0.19 | 1.07 | 0.67 | 0.27 | 0.25 | 0.25 | 0 | 0.58 | 2.98 | 0.67 | NA | 0.03 | 0.5 | NA | 1.55 | 0.82 | 0.93 | NA | 0.6 | 0.77
Results from laboratories 1.09, 1.11, 1.22, 1.25 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 284 (Certified Value 0.084 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.084855 | 0.0866 | 0.077 | 0.0819 | 0.0829 | 0.0828 | 0.085 | 0.084 | 0.089 | 0.08 | 0.08 | 0.082 | 0.089656 | 0.082 | 0.08 | 0.084 | 0.079 | 0.089 | 0.079 | 0.0846 | 0.0841 | 0.087 | 0.071 | 0.0873 |  |  | 1
| 1 | 2 | 0.085169 | 0.0857 | 0.077 | 0.0806 | 0.0835 | 0.0816 | 0.082 | 0.086 | 0.087 | 0.08 | 0.084 | 0.088 | 0.085295 | 0.083 | 0.0827 | 0.086 | 0.0797 | 0.083 | 0.084 | 0.0844 | 0.0795 | 0.087 | 0.071 | 0.0875 |  |  | 17
| 1 | 3 | 0.087661 | 0.0877 | 0.082 | 0.0807 | 0.0828 | 0.0809 | 0.086 | 0.083 | 0.086 | 0.078 | 0.085 | 0.088 | 0.088024 | 0.082 | 0.0832 | 0.085 | 0.0784 | 0.085 | 0.08 | 0.0837 | 0.0833 | 0.087 | 0.0713 | 0.0873 |  |  | 16
| 1 | 4 | 0.08785 | 0.0859 | 0.084 | 0.081 | 0.0836 | 0.0825 | 0.08 | 0.085 | 0.087 | 0.082 | 0.087 | 0.086 | 0.085258 | 0.081 | 0.0816 | 0.087 | 0.0811 | 0.085 | 0.084 | 0.0813 | 0.0821 | 0.087 | 0.077 | 0.0889 |  |  | 0.084071
| 1 | 5 | 0.085543 | 0.0868 | 0.089 | 0.0786 | 0.083 | 0.0812 | 0.084 | 0.086 | 0.087 | 0.081 | 0.085 | 0.086 | 0.088641 | 0.077 | 0.0819 | 0.085 | 0.0791 | 0.087 | 0.084 | 0.0861 | 0.0882 | 0.087 | 0.076 | 0.0873 |  |  | 48
| 1 | 6 | 0.085882 | 0.0877 | 0.082 | 0.0806 | 0.0836 | 0.0814 | 0.081 | 0.084 | 0.087 | 0.079 | 0.084 | 0.088 | 0.089222 | 0.081 | 0.0827 | 0.086 | 0.0792 | 0.087 | 0.084 | 0.0854 | 0.0822 | 0.087 | 0.077 | 0.0877
| Mean |  | 0.08616 | 0.086733 | 0.081833 | 0.080567 | 0.083233 | 0.081733 | 0.083 | 0.084667 | 0.087167 | 0.08 | 0.084167 | 0.086333 | 0.087682 | 0.081 | 0.082017 | 0.0855 | 0.079417 | 0.086 | 0.0825 | 0.08425 | 0.083233 | 0.087 | 0.073883 | 0.087667
| Median |  | 0.085713 | 0.0867 | 0.082 | 0.08065 | 0.08325 | 0.0815 | 0.083 | 0.0845 | 0.087 | 0.08 | 0.0845 | 0.087 | 0.088332 | 0.0815 | 0.0823 | 0.0855 | 0.07915 | 0.086 | 0.084 | 0.0845 | 0.08275 | 0.087 | 0.07365 | 0.0874
| Std Dev. |  | 0.0012847 | 0.00085479 | 0.004535 | 0.001082 | 0.00037238 | 0.00075277 | 0.0023664 | 0.0012111 | 0.00098319 | 0.0014142 | 0.0023166 | 0.0023381 | 0.0019432 | 0.0020976 | 0.0011479 | 0.0010488 | 0.00092394 | 0.0020976 | 0.0023452 | 0.001667 | 0.0028884 | 1.5202e-17 | 0.0030727 | 0.00062503
| Rel.Std.Dev. |  | 1.491% | 0.9855% | 5.542% | 1.343% | 0.4474% | 0.921% | 2.851% | 1.43% | 1.128% | 1.768% | 2.752% | 2.708% | 2.216% | 2.59% | 1.4% | 1.227% | 1.163% | 2.439% | 2.843% | 1.979% | 3.47% | 1.747e-14% | 4.159% | 0.713%
| PDM3 |  | 2.485% | 3.167% | -2.662% | -4.168% | -0.9964% | -2.781% | -1.274% | 0.7085% | 3.682% | -4.842% | 0.1137% | 2.691% | 4.296% | -3.653% | -2.444% | 1.7% | -5.536% | 2.294% | -1.869% | 0.2128% | -0.9964% | 3.484% | -12.12% | 4.277%
| Z-Score (Absolute) |  | 0.7 | 0.86 | 0.53 | 0.89 | 0.13 | 0.56 | 0.2 | 0.27 | 0.98 | 1.05 | 0.13 | 0.75 | 1.13 | 0.76 | 0.48 | 0.51 | 1.21 | 0.65 | 0.34 | 0.16 | 0.13 | 1.01 | 2.78 | 1.12
| Analytical results for Pb in OREAS 284 (Certified Value 23.6 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 24.1013 | 25.2 | 23 | 17.5909 | 22 | 21 | 21.99 | 22.9 | 27.1 | 22.9 | 21.6 | 23.5 | 24.7595 | 23.1 | 22.2 | 24 | 25.7 | 27.5 | 23.8 | 22 | 24.2 | 30 | 20.3 | 22.4 |  |  | 1
| 1 | 2 | 23.8087 | 25.5 | 22 | 12.7009 | 23 | 21 | 22.34 | 23.1 | 26.8 | 22.6 | 21.2 | 24.2 | 24.8205 | 21.6 | 22.1 | 24 | 25.1 | 26.1 | 25.7 | 22 | 24.3 | 29 | 17.5 | 22.7 |  |  | 18
| 1 | 3 | 25.0748 | 25.5 | 23 | 13.0041 | 22 | 23 | 23.64 | 22.6 | 25.6 | 23 | 22 | 23 | 24.2338 | 22.8 | 22.7 | 25 | 25.8 | 26 | 24.5 | 22 | 24.5 | 27 | 18.1 | 22.2 |  |  | 16
| 1 | 4 | 24.8103 | 25.5 | 24 | 14.8589 | 22 | 23 | 24.5 | 22.8 | 25.6 | 23.6 | 21.5 | 24.3 | 24.5747 | 20.6 | 23.5 | 25 | 26 | 25.7 | 25.4 | 21 | 24.3 | 31 | 16.6 | 23.1 |  |  | 23.5631
| 1 | 5 | 24.1167 | 25.6 | 20 | 15.8821 | 22 | 20 | 22.55 | 24 | 26.2 | 23.6 | 21.1 | 23.5 | 24.1197 | 20.1 | 23.7 | 24 | 25.6 | 25.9 | 25.7 | 23 | 24.1 | 29 | 16.3 | 22.3 |  |  | 49
| 1 | 6 | 24.2358 | 25.2 | 22 | 13.2249 | 22 | 19 | 22.31 | 23.3 | 25.1 | 23.1 | 21.5 | 23.3 | 24.4829 | 22.1 | 22.6 | 25 | 26 | 25.2 | 25.4 | 22 | 25 | 34 | 14.6 | 23.2
| Mean |  | 24.3579 | 25.4167 | 22.3333 | 14.5436 | 22.1667 | 21.1667 | 22.8883 | 23.1167 | 26.0667 | 23.1333 | 21.4833 | 23.6333 | 24.4985 | 21.7167 | 22.8 | 24.5 | 25.7 | 26.0667 | 25.0833 | 22 | 24.4 | 30 | 17.2333 | 22.65
| Median |  | 24.1762 | 25.5 | 22.5 | 14.0419 | 22 | 21 | 22.445 | 23 | 25.9 | 23.05 | 21.5 | 23.5 | 24.5288 | 21.85 | 22.65 | 24.5 | 25.75 | 25.95 | 25.4 | 22 | 24.3 | 29.5 | 17.05 | 22.55
| Std Dev. |  | 0.48148 | 0.17224 | 1.3663 | 1.933 | 0.40825 | 1.6021 | 0.9715 | 0.49565 | 0.77374 | 0.39833 | 0.31885 | 0.51251 | 0.27981 | 1.1923 | 0.66332 | 0.54772 | 0.33466 | 0.77115 | 0.7679 | 0.63246 | 0.32249 | 2.3664 | 1.9201 | 0.42308
| Rel.Std.Dev. |  | 1.977% | 0.6777% | 6.118% | 13.29% | 1.842% | 7.569% | 4.245% | 2.144% | 2.968% | 1.722% | 1.484% | 2.169% | 1.142% | 5.49% | 2.909% | 2.236% | 1.302% | 2.958% | 3.061% | 2.875% | 1.322% | 7.888% | 11.14% | 1.868%
| PDM3 |  | 3.373% | 7.866% | -5.219% | -38.28% | -5.926% | -10.17% | -2.864% | -1.895% | 10.62% | -1.824% | -8.826% | 0.2979% | 3.97% | -7.836% | -3.239% | 3.976% | 9.069% | 10.62% | 6.452% | -6.634% | 3.552% | 27.32% | -26.86% | -3.875%
| Z-Score (Absolute) |  | 0.6 | 1.12 | 0.39 | 4.21 | 0.47 | 0.96 | 0.12 | 0 | 1.44 | 0 | 0.81 | 0.25 | 0.67 | 0.69 | 0.16 | 0.67 | 1.26 | 1.44 | 0.96 | 0.55 | 0.63 | 3.37 | 2.89 | 0.23
| Analytical results for Pr in OREAS 284 (Certified Value 9.28 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.13; 1.12: 1.26; 1.13: 1.29 |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.03: 9.029; 1.04: 8.87; 1.09: 9.3; 1.10: 10.16; 1.11: 8.5; 1.12: 9.67; 1.13: 7 |  |  | 1
| 1 | 2 | 1.03: 8.7612; 1.04: 8.97; 1.09: 9.4; 1.10: 10.254; 1.11: 9.9; 1.12: 9.43; 1.13: 8 |  |  | 4
| 1 | 3 | 1.03: 9.3006; 1.04: 8.86; 1.09: 9.6; 1.10: 10.3614; 1.11: 9.2; 1.12: 9.49; 1.13: 7 |  |  | 16
| 1 | 4 | 1.03: 8.9665; 1.04: 8.7; 1.09: 9.1; 1.10: 9.5533; 1.11: 9.2; 1.12: 9.77; 1.13: 7 |  |  | 9.2837
| 1 | 5 | 1.03: 9.1177; 1.04: 8.92; 1.09: 8.5; 1.10: 9.413; 1.11: 8.7; 1.12: 9.78; 1.13: 7 |  |  | 50
| 1 | 6 | 1.03: 9.0866; 1.04: 8.89; 1.09: 9.7; 1.10: 9.5498; 1.11: 8.2; 1.12: 10.01; 1.13: 7
| Mean |  | 1.03: 9.0436; 1.04: 8.8683; 1.09: 9.2667; 1.10: 9.8819; 1.11: 8.95; 1.12: 9.6917; 1.13: 7.1667
| Median |  | 1.03: 9.0578; 1.04: 8.88; 1.09: 9.35; 1.10: 9.8567; 1.11: 8.95; 1.12: 9.72; 1.13: 7
| Std Dev. |  | 1.03: 0.17849; 1.04: 0.091524; 1.09: 0.43205; 1.10: 0.42044; 1.11: 0.6091; 1.12: 0.21208; 1.13: 0.40825
| Rel.Std.Dev. |  | 1.03: 1.974%; 1.04: 1.032%; 1.09: 4.662%; 1.10: 4.255%; 1.11: 6.806%; 1.12: 2.188%; 1.13: 5.696%
| PDM3 |  | 1.03: -2.586%; 1.04: -4.474%; 1.09: -0.1834%; 1.10: 6.444%; 1.11: -3.594%; 1.12: 4.394%; 1.13: -22.8%
| Z-Score (Absolute) |  | 1.03: 0.31; 1.04: 0.79; 1.09: 0.31; 1.10: 1.99; 1.11: 0.56; 1.12: 1.47; 1.13: NA
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 284 (Certified Value 170 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 0
0
| 1 | 1 | 162.4791 | 180.15 | 156 | 173 | 112.1 | 163 | 152 | 167.5 | 155.5 | 164 | 177.1292 | 154.7 | 168.2 | 168 | 184 | 171.2 | 165 | 184.94 | 171.44 | 114 | 150.9 | [5 empty] | 1
| 1 | 2 | 160.9159 | 178.39 | 160 | 176.4 | 184.4 | 171.5 | 157 | 155 | 161 | 166.5 | 175.9447 | 142.6 | 172 | 168 | 179 | 164.3 | 173 | 183.55 | 173.66 | 112 | 148.3 | [5 empty] | 34
| 1 | 3 | 163.163 | 181.17 | 168 | 179 | 135.6 | 163 | 155.5 | 158 | 164.5 | 162 | 175.8488 | 165.3 | 172.9 | 167 | 182 | 165.9 | 166 | 184.21 | 172.51 | 122 | 149.3 | [5 empty] | 16
| 1 | 4 | 167.9547 | 180.34 | 184 | 183.9 | 116.7 | 161 | 155 | 168.5 | 161 | 172.5 | 177.5682 | 166.3 | 174.5 | 170 | 188 | 164.5 | 176 | 183.68 | 173.39 | 116 | 153.6 | [5 empty] | 169.6048
| 1 | 5 | 162.8792 | 179.6 | 174 | 180.2 | 121.4 | 173 | 160.5 | 162 | 152.5 | 163 | 178.6081 | 165.1 | 173.2 | 167 | 189 | 173.3 | 181 | 190.07 | 175.37 | 106 | 154.1 | [5 empty] | 51
| 1 | 6 | 169.7952 | 179.05 | 188 | 184.1 | 107.5 | 167 | 162.5 | 162 | 163 | 166 | 179.0218 | 172.7 | 176.1 | 162 | 186 | 166.1 | 175 | 181.19 | 175.96 | 114 | 150.1
| Mean |  | 164.5312 | 179.7833 | 171.6667 | 179.4333 | 129.6167 | 166.4167 | 157.0833 | 162.1667 | 159.5833 | 165.6667 | 177.3535 | 161.1167 | 172.8167 | 167 | 184.6667 | 167.55 | 172.6667 | 184.6067 | 173.7217 | 114 | 151.05
| Median |  | 163.0211 | 179.875 | 171 | 179.6 | 119.05 | 165 | 156.25 | 162 | 161 | 165 | 177.3487 | 165.2 | 173.05 | 167.5 | 185 | 166 | 174 | 183.945 | 173.525 | 114 | 150.5
| Std Dev. |  | 3.5018 | 0.98792 | 12.801 | 4.3149 | 28.5214 | 4.9439 | 3.8395 | 5.2409 | 4.6197 | 3.7639 | 1.3193 | 10.7559 | 2.6739 | 2.6833 | 3.7771 | 3.7703 | 6.1536 | 2.9586 | 1.7038 | 5.2154 | 2.339
| Rel.Std.Dev. |  | 2.128% | 0.5495% | 7.457% | 2.405% | 22% | 2.971% | 2.444% | 3.232% | 2.895% | 2.272% | 0.7439% | 6.676% | 1.547% | 1.607% | 2.045% | 2.25% | 3.564% | 1.603% | 0.9808% | 4.575% | 1.549%
| PDM3 |  | -2.991% | 6.001% | 1.216% | 5.795% | -23.58% | -1.88% | -7.383% | -4.386% | -5.909% | -2.322% | 4.569% | -5.005% | 1.894% | -1.536% | 8.881% | -1.211% | 1.805% | 8.845% | 2.427% | -32.78% | -10.94%
| Z-Score (Absolute) |  | 0.25 | 1.28 | 0.47 | 1.25 | 3.75 | 0.06 | 0.99 | 0.48 | 0.74 | 0.13 | 1.04 | 0.59 | 0.58 | 0 | 1.77 | 0.06 | 0.57 | 1.77 | 0.67 | 5.32 | 1.6
| Analytical results for Re in OREAS 284 (Indicative Value 0.002 ppm) | [27 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.24 | 1.26 | 1.31 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 3
3
| 1 | 1 | < 0.01 | < 0.002 | 0.008 | 0.013 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.006 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.002 | [9 empty] | 1
| 1 | 2 | < 0.01 | < 0.002 | 0.007 | 0.013 | 0.003 | < 0.002 | 0.002 | 0.003 | 0.003 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | < 0.002 | [9 empty] | 2
| 1 | 3 | < 0.01 | < 0.002 | 0.008 | 0.012 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | < 0.002 | [9 empty] | 16
| 1 | 4 | < 0.01 | < 0.002 | 0.007 | 0.013 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | < 0.002 | [9 empty] | 0.0015
| 1 | 5 | < 0.01 | < 0.002 | 0.008 | 0.009 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | 0.002 | [9 empty] | 8
| 1 | 6 | < 0.01 | < 0.002 | 0.005 | 0.011 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | 0.002
| Mean |  | 1.09: 0.0071667; 1.12: 0.011833; 1.13: 0.003; 1.14: 0.002; 1.15: 0.002; 1.16: 0.0025; 1.17: 0.003; 1.18: 0.002; 1.24: 0.003; 1.25: 0.002
| Median |  | 1.09: 0.0075; 1.12: 0.0125; 1.13: 0.003; 1.14: 0.002; 1.15: 0.002; 1.16: 0.0025; 1.17: 0.002; 1.18: 0.002; 1.24: 0.003; 1.25: 0.002
| Std Dev. |  | 1.09: 0.001169; 1.12: 0.0016021; 1.16: 0.00070711; 1.17: 0.0017321; 1.18: 0; 1.25: 0
| Rel.Std.Dev. |  | 1.09: 16.31%; 1.12: 13.54%; 1.16: 28.28%; 1.17: 57.74%; 1.18: 0%; 1.25: 0%
| PDM3 |  | 1.09: 377.8%; 1.12: 688.9%; 1.13: 100%; 1.14: 33.33%; 1.15: 33.33%; 1.16: 66.67%; 1.17: 100%; 1.18: 33.33%; 1.24: 100%; 1.25: 33.33%
| Z-Score (Absolute) |  | 2.43 | 0.81 | 4.18 | 7.96 | 0.54 | 0.67 | 0.67 | 0.4 | 0.54 | 0 | 18.61 | 0.4 | 18.61 | 0.81 | 0.4 | 0.54 | 0.4
| Analytical results for S in OREAS 284 (Certified Value 0.262 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.23782 | 0.26438 | 0.2566 | 0.24 | 0.3351 | 0.31 | 0.2433 | 0.26 | 0.26 | 0.28 | 0.25 | 0.26 | 0.26 | 0.2608 | 0.24 | 0.26 | 0.255 | 0.26 | 0.2 | 0.252 | 0.2695 | 0.26 | 0.291 | 0.25 | 0.26 |  | 1
| 1 | 2 | 0.23294 | 0.27561 | 0.2558 | 0.23 | 0.3089 | 0.31 | 0.2418 | 0.26 | 0.26 | 0.27 | 0.25 | 0.26 | 0.28 | 0.26384 | 0.24 | 0.26 | 0.262 | 0.27 | 0.2 | 0.268 | 0.2698 | 0.28 | 0.292 | 0.25 | 0.26 |  | 11
| 1 | 3 | 0.23433 | 0.26995 | 0.2611 | 0.24 | 0.3038 | 0.31 | 0.241 | 0.27 | 0.26 | 0.27 | 0.25 | 0.27 | 0.27 | 0.26099 | 0.24 | 0.26 | 0.256 | 0.26 | 0.2 | 0.257 | 0.2656 | 0.28 | 0.263 | 0.24 | 0.25 |  | 16
| 1 | 4 | 0.23407 | 0.27014 | 0.2571 | 0.26 | 0.317 | 0.31 | 0.245 | 0.25 | 0.26 | 0.27 | 0.26 | 0.27 | 0.27 | 0.26097 | 0.23 | 0.27 | 0.259 | 0.27 | 0.2 | 0.266 | 0.2584 | 0.28 | 0.268 | 0.24 | 0.26 |  | 0.26173
| 1 | 5 | 0.23428 | 0.26263 | 0.253 | 0.26 | 0.3071 | 0.31 | 0.2387 | 0.27 | 0.27 | 0.27 | 0.26 | 0.27 | 0.27 | 0.25911 | 0.22 | 0.27 | 0.256 | 0.26 | 0.2 | 0.267 | 0.2713 | 0.27 | 0.283 | 0.25 | 0.26 |  | 52
| 1 | 6 | 0.23197 | 0.27134 | 0.2621 | 0.25 | 0.3184 | 0.31 | 0.2394 | 0.25 | 0.26 | 0.27 | 0.25 | 0.27 | 0.27 | 0.26307 | 0.23 | 0.27 | 0.26 | 0.26 | 0.2 | 0.272 | 0.2713 | 0.26 | 0.286 | 0.25 | 0.25
| Mean |  | 0.23424 | 0.26901 | 0.25762 | 0.24667 | 0.31505 | 0.31 | 0.24153 | 0.26 | 0.26167 | 0.27167 | 0.25333 | 0.26667 | 0.27 | 0.26146 | 0.23333 | 0.265 | 0.258 | 0.26333 | 0.2 | 0.26367 | 0.26765 | 0.27167 | 0.2805 | 0.24667 | 0.25667
| Median |  | 0.23418 | 0.27005 | 0.25685 | 0.245 | 0.31295 | 0.31 | 0.2414 | 0.26 | 0.26 | 0.27 | 0.25 | 0.27 | 0.27 | 0.26098 | 0.235 | 0.265 | 0.2575 | 0.26 | 0.2 | 0.2665 | 0.26965 | 0.275 | 0.2845 | 0.25 | 0.26
| Std Dev. |  | 0.0019836 | 0.0047601 | 0.0034102 | 0.012111 | 0.011354 | 0 | 0.0023695 | 0.0089443 | 0.0040825 | 0.0040825 | 0.005164 | 0.005164 | 0.0063246 | 0.001711 | 0.008165 | 0.0054772 | 0.0027568 | 0.005164 | 3.0405e-17 | 0.0075542 | 0.0049891 | 0.0098319 | 0.012178 | 0.005164 | 0.005164
| Rel.Std.Dev. |  | 0.8468% | 1.769% | 1.324% | 4.91% | 3.604% | 0% | 0.981% | 3.44% | 1.56% | 1.503% | 2.038% | 1.936% | 2.342% | 0.6544% | 3.499% | 2.067% | 1.069% | 1.961% | 1.52e-14% | 2.865% | 1.864% | 3.619% | 4.341% | 2.094% | 2.012%
| PDM3 |  | -10.51% | 2.781% | -1.572% | -5.756% | 20.37% | 18.44% | -7.717% | -0.6614% | -0.02463% | 3.796% | -3.209% | 1.886% | 3.159% | -0.102% | -10.85% | 1.249% | -1.426% | 0.6122% | -23.59% | 0.7395% | 2.261% | 3.796% | 7.171% | -5.756% | -1.935%
| Z-Score (Absolute) |  | 2.72 | 0.63 | 0.47 | 1.52 | 5.06 | 4.57 | 2.02 | 0.24 | 0.08 | 0.88 | 0.88 | 0.4 | 0.72 | 0.1 | 2.81 | 0.24 | 0.43 | 0.08 | NA | 0.11 | 0.5 | 0.88 | 1.73 | 1.52 | 0.56
Results from laboratory 1.24 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 284 (Certified Value 0.47 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | < 2.5 | 0.45 | 0.4 | < 5 | 0.48 | 0.37 | 0.43 | 0.44 | 0.44 | 0.44 | 0.5 | 0.4576 | 0.6 | 0.46 | 0.5 | 0.49 | 0.4 | 0.45 | 0.46 | 0.45 | 4.98 | 0.98 | 0.49 |  |  |  | 1
| 1 | 2 | < 2.5 | 0.46 | 0.3 | < 5 | 0.43 | 0.46 | 0.47 | 0.43 | 0.41 | 0.43 | 0.49 | 0.52575 | 0.6 | 0.46 | 0.5 | 0.44 | 0.4 | 0.48 | 0.5 | 0.47 | 4.13 | 0.77 | 0.52 |  |  |  | 36
| 1 | 3 | < 2.5 | 0.48 | 0.2 | < 5 | 0.47 | 0.48 | 0.46 | 0.43 | 0.43 | 0.53 | 0.51 | 0.48141 | 0.8 | 0.45 | 0.4 | 0.48 | 0.4 | 0.47 | 0.48 | 0.55 |  | 0.75 | 0.48 |  |  |  | 16
| 1 | 4 | < 2.5 | 0.46 | 0.2 | < 5 | 0.47 | 0.44 | 0.44 | 0.42 | 0.47 | 0.43 | 0.52 | 0.50555 | 0.7 | 0.48 | 0.4 | 0.49 | 0.4 | 0.47 | 0.51 | 0.46 | 5.48 | 0.65 | 0.5 |  |  |  | 0.46625
| 1 | 5 | < 2.5 | 0.51 | 0.4 | < 5 | 0.44 | 0.4 | 0.45 | 0.43 | 0.5 | 0.44 | 0.46 | 0.45084 | 0.5 | 0.49 | 0.4 | 0.51 | 0.5 | 0.48 | 0.48 | 0.49 | 5.23 | 0.7 | 0.5 |  |  |  | 53
| 1 | 6 | < 2.5 | 0.46 | 0.2 | < 5 | 0.44 | 0.37 | 0.42 | 0.43 | 0.45 | 0.48 | 0.51 | 0.5176 | 0.7 | 0.46 | 0.5 | 0.49 | 0.4 | 0.47 | 0.46 | 0.46 |  | 0.71 | 0.5
| Mean |  |  | 0.47 | 0.28333 |  | 0.455 | 0.42 | 0.445 | 0.43 | 0.45 | 0.45833 | 0.49833 | 0.48979 | 0.65 | 0.46667 | 0.45 | 0.48333 | 0.41667 | 0.47 | 0.48167 | 0.48 | 4.955 | 0.76 | 0.49833
| Median |  |  | 0.46 | 0.25 |  | 0.455 | 0.42 | 0.445 | 0.43 | 0.445 | 0.44 | 0.505 | 0.49348 | 0.65 | 0.46 | 0.45 | 0.49 | 0.4 | 0.47 | 0.48 | 0.465 | 5.105 | 0.73 | 0.5
| Std Dev. |  |  | 0.021909 | 0.098319 |  | 0.020736 | 0.046904 | 0.018708 | 0.0063246 | 0.031623 | 0.039707 | 0.02137 | 0.031427 | 0.10488 | 0.015055 | 0.054772 | 0.023381 | 0.040825 | 0.010954 | 0.020412 | 0.036878 | 0.58666 | 0.11559 | 0.013292
| Rel.Std.Dev. |  |  | 4.661% | 34.7% |  | 4.557% | 11.17% | 4.204% | 1.471% | 7.027% | 8.663% | 4.288% | 6.416% | 16.14% | 3.226% | 12.17% | 4.837% | 9.798% | 2.331% | 4.238% | 7.683% | 11.84% | 15.21% | 2.667%
| PDM3 |  |  | 0.8037% | -39.23% |  | -2.413% | -9.92% | -4.558% | -7.775% | -3.486% | -1.699% | 6.88% | 5.049% | 39.41% | 0.08875% | -3.486% | 3.663% | -10.64% | 0.8037% | 3.306% | 2.948% | 962.7% | 63% | 6.88%
| Z-Score (Absolute) |  | 23.96 | 0.31 | NA | 62.87 | 0.78 | 1.87 | 1.09 | 1.56 | 0.93 | 0.67 | 0.57 | 0.3 | NA | 0.41 | NA | 0.1 | NA | 0.31 | 0.05 | 0 | 139.27 | 8.71 | 0.57
Results from laboratories 1.09, 1.20, 1.22 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 284 (Certified Value 9.03 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 8.6718 | 9.2 | 9 | 10.7015 | 9 | 8.39 | 8.6 | 8.9 | 9.2 | 8.5 | 8 | 9.3 | 9.0489 | 9.8 | 8.5 | 8 | 9.6 | 9 | 9.3 | 9.1 | 8.5 | 9 | 8.4 | 7.4 |  |  | 1
| 1 | 2 | 8.6929 | 9.2 | 9 | 10.5307 | 9 | 8.38 | 8.6 | 9.3 | 9 | 7.8 | 8.6 | 9.5 | 9.2288 | 10.1 | 8.7 | 8 | 9.4 | 8 | 9.9 | 9 | 8.5 | 9 | 8.3 | 7 |  |  | 37
| 1 | 3 | 8.8028 | 9 | 9 | 10.8936 | 9 | 8.34 | 8.9 | 8.8 | 8.6 | 8.2 | 8.5 | 9.4 | 8.9897 | 9.5 | 8.8 | 8 | 9.8 | 8 | 9.3 | 9 | 8.8 | 9 | 8.48 | 6.9 |  |  | 16
| 1 | 4 | 8.9069 | 9.3 | 9 | 10.2842 | 9 | 8.36 | 8.4 | 8.8 | 9.3 | 8.4 | 8.4 | 9.8 | 9.1898 | 9.7 | 8.7 | 8 | 10 | 8 | 9.9 | 8.8 | 8.8 | 9 | 9.1 | 7.3 |  |  | 9.0308
| 1 | 5 | 8.6564 | 9.1 | 9 | 10.0373 | 9 | 8.27 | 8.7 | 9.3 | 8.8 | 8.4 | 8.3 | 9.4 | 8.9892 | 8.9 | 8.9 | 8 | 9.8 | 9 | 10.1 | 9.2 | 8.8 | 9 | 8.8 | 6.8 |  |  | 54
| 1 | 6 | 8.6992 | 9.1 | 9 | 10.0985 | 9 | 8.36 | 8.2 | 9 | 9 | 8.2 | 8.7 | 9.7 | 9.1828 | 9.4 | 8.8 | 8 | 9.7 | 8 | 9.6 | 9 | 9 | 9 | 9 | 6.8
| Mean |  | 8.7383 | 9.15 | 9 | 10.4243 | 9 | 8.35 | 8.5667 | 9.0167 | 8.9833 | 8.25 | 8.4167 | 9.5167 | 9.1049 | 9.5667 | 8.7333 | 8 | 9.7167 | 8.3333 | 9.6833 | 9.0167 | 8.7333 | 9 | 8.68 | 7.0333
| Median |  | 8.696 | 9.15 | 9 | 10.4075 | 9 | 8.36 | 8.6 | 8.95 | 9 | 8.3 | 8.45 | 9.45 | 9.1158 | 9.6 | 8.75 | 8 | 9.75 | 8 | 9.75 | 9 | 8.8 | 9 | 8.64 | 6.95
| Std Dev. |  | 0.09726 | 0.10488 | 0 | 0.34174 | 0 | 0.042895 | 0.24221 | 0.23166 | 0.25626 | 0.251 | 0.24833 | 0.19408 | 0.10812 | 0.40825 | 0.13663 | 0 | 0.20412 | 0.5164 | 0.33714 | 0.13292 | 0.19664 | 0 | 0.33347 | 0.2582
| Rel.Std.Dev. |  | 1.113% | 1.146% | 0% | 3.278% | 0% | 0.5137% | 2.827% | 2.569% | 2.853% | 3.042% | 2.95% | 2.039% | 1.187% | 4.267% | 1.564% | 0% | 2.101% | 6.197% | 3.482% | 1.474% | 2.252% | 0% | 3.842% | 3.671%
| PDM3 |  | -3.238% | 1.32% | -0.3406% | 15.43% | -0.3406% | -7.538% | -5.139% | -0.156% | -0.5251% | -8.646% | -6.8% | 5.381% | 0.8206% | 5.934% | -3.293% | -11.41% | 7.595% | -7.723% | 7.226% | -0.156% | -3.293% | -0.3406% | -3.884% | -22.12%
| Z-Score (Absolute) |  | 0.4 | 0.27 | NA | 2.33 | NA | 1.02 | 0.67 | 0.05 | 0 | 1.19 | 0.92 | 0.86 | 0.2 | 0.94 | 0.4 | NA | 1.19 | NA | 1.13 | 0.05 | 0.4 | NA | 0.49 | 3.16
Results from laboratories 1.09, 1.11, 1.22, 1.24 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 284 (Indicative Value 0.7 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | < 1.5 | 0.7 | 0.3 | < 0.1 | < 5 | 0.39 | 1.2 | 1 | 1 | 1 | < 1 | < 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 | 3 | < 2 | [5 empty] | 1
| 1 | 2 | < 1.5 | 0.9 | 0.3 | < 0.1 | < 5 | 0.4 | < 0.3 | 1 | 1 | 1 | < 1 | < 1 | < 2 | < 2 | < 1 | 0.5 | < 1 | < 1 | < 0.5 | 4 | < 2 | [5 empty] | 38
| 1 | 3 | < 1.5 | 1 | 0.3 | < 0.1 | < 5 | 0.43 | 0.8 | 1 | 1 | 1 | < 1 | < 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 | 3.4 | < 2 | [5 empty] | 16
| 1 | 4 | < 1.5 | 0.9 | 0.6 | < 0.1 | < 5 | 0.42 | 0.8 | 1 | 1 | < 1 | < 1 | 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 | 4 | < 2 | [5 empty] | 0.70278
| 1 | 5 | < 1.5 | 1.1 | 0.7 | < 0.1 | < 5 | 0.43 | 0.5 | < 1 | 1 | < 1 | < 1 | < 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 | 3 | < 2 | [5 empty] | 7
| 1 | 6 | < 1.5 | 0.9 | 0.8 | < 0.1 | < 5 | 0.43 | 1 | 1 | 1 | 1 | < 1 | 1 | < 2 | < 2 | < 1 | 0.5 | < 1 | < 1 | < 0.5 | 3 | < 2
| Mean |  | 1.04: 0.91667; 1.09: 0.5; 1.12: 0.41667; 1.13: 0.86; 1.14: 1; 1.15: 1; 1.16: 1; 1.18: 1; 1.23: 0.5; 1.31: 3.4
| Median |  | 1.04: 0.9; 1.09: 0.45; 1.12: 0.425; 1.13: 0.8; 1.14: 1; 1.15: 1; 1.16: 1; 1.18: 1; 1.23: 0.5; 1.31: 3.2
| Std Dev. |  | 1.04: 0.13292; 1.09: 0.22804; 1.12: 0.017512; 1.13: 0.26077; 1.14: 0; 1.15: 0; 1.16: 0; 1.18: 0; 1.23: 0; 1.31: 0.4899
| Rel.Std.Dev. |  | 1.04: 14.5%; 1.09: 45.61%; 1.12: 4.203%; 1.13: 30.32%; 1.14: 0%; 1.15: 0%; 1.16: 0%; 1.18: 0%; 1.23: 0%; 1.31: 14.41%
| PDM3 |  | 1.04: 30.43%; 1.09: -28.85%; 1.12: -40.71%; 1.13: 22.37%; 1.14: 42.29%; 1.15: 42.29%; 1.16: 42.29%; 1.18: 42.29%; 1.23: -28.85%; 1.31: 383.8%
| Z-Score (Absolute) |  | 0.02 | 0.49 | 0.67 | 1.93 | 4.91 | 0.91 | 0 | 0.49 | 0.72 | 0.26 | 0.67 | 0.21 | 0.72 | 0.72 | 0.67 | 1.14 | 0.67 | 0.67 | 1.37 | 7.42 | 0.72
| Analytical results for Sm in OREAS 284 (Certified Value 7.27 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab | [5 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.13; 1.12: 1.26; 1.13: 1.29 | [5 empty] | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.03: 7.1488; 1.04: 7.35; 1.09: 6.8; 1.10: 7.3041; 1.11: 6.4; 1.12: 7.68; 1.13: 5 | [5 empty] | 1
| 1 | 2 | 1.03: 7.2034; 1.04: 7.62; 1.09: 7.3; 1.10: 7.1034; 1.11: 8.5; 1.12: 7.91; 1.13: 5 | [5 empty] | 4
| 1 | 3 | 1.03: 7.3214; 1.04: 7.51; 1.09: 7.8; 1.10: 7.302; 1.11: 7.3; 1.12: 7.55; 1.13: 5 | [5 empty] | 16
| 1 | 4 | 1.03: 7.2951; 1.04: 7.35; 1.09: 7.3; 1.10: 6.7116; 1.11: 7.1; 1.12: 7.67; 1.13: 5 | [5 empty] | 7.2704
| 1 | 5 | 1.03: 7.2229; 1.04: 7.51; 1.09: 6.7; 1.10: 6.3773; 1.11: 6.6; 1.12: 7.69; 1.13: 5 | [5 empty] | 55
| 1 | 6 | 1.03: 7.2034; 1.04: 7.44; 1.09: 7.6; 1.10: 6.7326; 1.11: 6.3; 1.12: 7.83; 1.13: 5
| Mean |  | 1.03: 7.2325; 1.04: 7.4633; 1.09: 7.25; 1.10: 6.9218; 1.11: 7.0333; 1.12: 7.7217; 1.13: 5
| Median |  | 1.03: 7.2132; 1.04: 7.475; 1.09: 7.3; 1.10: 6.918; 1.11: 6.85; 1.12: 7.685; 1.13: 5
| Std Dev. |  | 1.03: 0.064209; 1.04: 0.10501; 1.09: 0.43243; 1.10: 0.3742; 1.11: 0.81894; 1.12: 0.12813; 1.13: 0
| Rel.Std.Dev. |  | 1.03: 0.8878%; 1.04: 1.407%; 1.09: 5.965%; 1.10: 5.406%; 1.11: 11.64%; 1.12: 1.659%; 1.13: 0%
| PDM3 |  | 1.03: -0.5219%; 1.04: 2.653%; 1.09: -0.2812%; 1.10: -4.795%; 1.11: -3.261%; 1.12: 6.206%; 1.13: -31.23%
| Z-Score (Absolute) |  | 1.03: 0.03; 1.04: 0.7; 1.09: 0.03; 1.10: 1; 1.11: 0.65; 1.12: 1.51; 1.13: NA
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 284 (Certified Value 4.79 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 4.7566 | 5 | 3 | < 10 | 4.94 | 4.4 | 4.5 | 4.6 | 4.8 | 4.5 | 4.8 | 5.1174 | 4.9 | 4.5 | 4.6 | 4.9 | 4.7 | 4.4 | 5.3 | 4.7 | 12 | 3.4 | 4.8 |  |  |  | 1
| 1 | 2 | 4.7781 | 4.9 | 2 | < 10 | 4.64 | 4.6 | 4.6 | 4.5 | 4.6 | 4.7 | 4.9 | 5.1254 | 5 | 4.9 | 4.6 | 4.7 | 4.5 | 4.8 | 5.2 | 4.9 | 14 | 3.4 | 4.8 |  |  |  | 13
| 1 | 3 | 4.7521 | 5.1 | 3 | < 10 | 3.94 | 4.5 | 4.5 | 4.4 | 4.8 | 4.8 | 4.8 | 4.8473 | 5 | 4.8 | 4.8 | 4.7 | 4.5 | 4.6 | 5.3 | 4.7 | 11 | 3.4 | 4.8 |  |  |  | 16
| 1 | 4 | 4.9244 | 4.9 | 4 | < 10 | 4.74 | 7.3 | 4.6 | 4.7 | 5.2 | 4.7 | 5.1 | 4.9091 | 5.1 | 5.1 | 4.6 | 5 | 4.5 | 4.8 | 5.2 | 5 | 12 | 3.5 | 4.9 |  |  |  | 4.7877
| 1 | 5 | 4.8094 | 4.9 | 5 | < 10 | 4.64 | 4.2 | 4.7 | 4.6 | 4.9 | 4.5 | 4.8 | 5.0312 | 4.7 | 4.8 | 4.6 | 4.9 | 4.6 | 4.9 | 5.4 | 4.8 | 12 | 3.5 | 4.8 |  |  |  | 56
| 1 | 6 | 4.8749 | 4.9 | 4 | < 10 | 3.54 | 4 | 4.6 | 4.8 | 4.7 | 4.7 | 5 | 5.1679 | 4.9 | 4.6 | 4.9 | 4.9 | 4.5 | 4.8 | 5.4 | 5 | 13 | 3.4 | 4.9
| Mean |  | 4.8159 | 4.95 | 3.5 |  | 4.4067 | 4.8333 | 4.5833 | 4.6 | 4.8333 | 4.65 | 4.9 | 5.0331 | 4.9333 | 4.7833 | 4.6833 | 4.85 | 4.55 | 4.7167 | 5.3 | 4.85 | 12.3333 | 3.4333 | 4.8333
| Median |  | 4.7938 | 4.9 | 3.5 |  | 4.64 | 4.45 | 4.6 | 4.6 | 4.8 | 4.7 | 4.85 | 5.0743 | 4.95 | 4.8 | 4.6 | 4.9 | 4.5 | 4.8 | 5.3 | 4.85 | 12 | 3.4 | 4.8
| Std Dev. |  | 0.069736 | 0.083666 | 1.0488 |  | 0.54283 | 1.2275 | 0.075277 | 0.14142 | 0.20656 | 0.12247 | 0.12649 | 0.12937 | 0.13663 | 0.2137 | 0.13292 | 0.12247 | 0.083666 | 0.18348 | 0.089443 | 0.13784 | 1.0328 | 0.05164 | 0.05164
| Rel.Std.Dev. |  | 1.448% | 1.69% | 29.97% |  | 12.32% | 25.4% | 1.642% | 3.074% | 4.274% | 2.634% | 2.581% | 2.57% | 2.769% | 4.468% | 2.838% | 2.525% | 1.839% | 3.89% | 1.688% | 2.842% | 8.374% | 1.504% | 1.068%
| PDM3 |  | 0.5907% | 3.391% | -26.9% |  | -7.958% | 0.9538% | -4.268% | -3.92% | 0.9538% | -2.875% | 2.346% | 5.126% | 3.043% | -0.09052% | -2.179% | 1.302% | -4.964% | -1.483% | 10.7% | 1.302% | 157.6% | -28.29% | 0.9538%
| Z-Score (Absolute) |  | 0.09 | 0.59 | NA | 0.84 | 2.16 | 0 | 1.26 | 1.18 | 0 | 0.93 | 0.34 | 1.01 | 0.51 | 0.25 | 0.76 | 0.08 | 1.43 | 0.59 | 2.36 | 0.08 | 37.93 | 7.08 | 0
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 284 (Certified Value 161 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
0
| 1 | 1 | 166.0524 | 152.9854 | 166.15 | 161 | 157.0405 | 167 | 176 | 164 | 159.5 | 165.5 | 158 | 148 | 162 | 166.0731 | 152 | 154 | 160 | 169 | 169 | 151 | 158 | 155.07 | 117.99 | 165 | 106 | 147.8 | 1
| 1 | 2 | 164.9618 | 156.325 | 166.12 | 161 | 156.0652 | 166 | 176 | 168 | 161.5 | 162 | 155.5 | 155 | 171 | 163.2644 | 151 | 155.3 | 161 | 170 | 157 | 160 | 157 | 159.86 | 127.42 | 167 | 107 | 148.4 | 14
| 1 | 3 | 166.5592 | 159.372 | 166.25 | 169 | 157.4898 | 164 | 176 | 170 | 157.5 | 160 | 154.5 | 154 | 164.5 | 163.6349 | 150 | 158 | 162 | 167 | 161 | 153 | 155 | 159.59 | 107.78 | 160 | 105 | 148.2 | 16
| 1 | 4 | 163.5702 | 158.9646 | 166.79 | 166 | 158.43 | 168 | 178 | 161 | 160 | 164 | 161 | 158.5 | 168 | 166.0674 | 148 | 156.3 | 160 | 174 | 157 | 160 | 154 | 158.27 |  | 160 | 111 | 154.9 | 161.3172
| 1 | 5 | 165.3734 | 154.6756 | 166.78 | 160 | 152.451 | 166 | 174 | 168 | 165 | 162.5 | 160.5 | 152 | 166.5 | 165.6329 | 142 | 158.3 | 162 | 175 | 163 | 160 | 160 | 160.6 | 116.17 | 166 | 109 | 149.6 | 57
| 1 | 6 | 165.932 | 157.4433 | 166.71 | 172 | 156.5358 | 168 | 175 | 158 | 161.5 | 168 | 154.5 | 154.5 | 169.5 | 166.241 | 148 | 159.9 | 164 | 171 | 158 | 160 | 157 | 160.45 | 116.36 | 172 | 112 | 151
| Mean |  | 165.4082 | 156.6276 | 166.4667 | 164.8333 | 156.3354 | 166.5 | 175.8333 | 164.8333 | 160.8333 | 163.6667 | 157.3333 | 153.6667 | 166.9167 | 165.1523 | 148.5 | 156.9667 | 161.5 | 171 | 160.8333 | 157.3333 | 156.8333 | 158.9733 | 117.144 | 165 | 108.3333 | 149.9833
| Median |  | 165.6527 | 156.8841 | 166.48 | 163.5 | 156.7882 | 166.5 | 176 | 166 | 160.75 | 163.25 | 156.75 | 154.25 | 167.25 | 165.8502 | 149 | 157.15 | 161.5 | 170.5 | 159.5 | 160 | 157 | 159.725 | 116.36 | 165.5 | 108 | 149
| Std Dev. |  | 1.0576 | 2.483 | 0.32537 | 4.9565 | 2.0697 | 1.5166 | 1.3292 | 4.6655 | 2.5232 | 2.8225 | 2.9439 | 3.4881 | 3.3078 | 1.3393 | 3.5637 | 2.1667 | 1.5166 | 3.0332 | 4.6655 | 4.1793 | 2.137 | 2.0844 | 6.9921 | 4.5607 | 2.8048 | 2.6761
| Rel.Std.Dev. |  | 0.6394% | 1.585% | 0.1955% | 3.007% | 1.324% | 0.9109% | 0.7559% | 2.83% | 1.569% | 1.725% | 1.871% | 2.27% | 1.982% | 0.8109% | 2.4% | 1.38% | 0.9391% | 1.774% | 2.901% | 2.656% | 1.363% | 1.311% | 5.969% | 2.764% | 2.589% | 1.784%
| PDM3 |  | 2.536% | -2.907% | 3.192% | 2.18% | -3.088% | 3.213% | 8.998% | 2.18% | -0.3% | 1.456% | -2.47% | -4.743% | 3.471% | 2.377% | -7.945% | -2.697% | 0.1133% | 6.002% | -0.3% | -2.47% | -2.78% | -1.453% | -27.38% | 2.283% | -32.84% | -7.026%
| Z-Score (Absolute) |  | 0.72 | 0.67 | 0.89 | 0.63 | 0.71 | 0.9 | 2.37 | 0.63 | 0 | 0.45 | 0.55 | 1.13 | 0.96 | 0.68 | 1.95 | 0.61 | 0.11 | 1.61 | 0 | 0.55 | 0.63 | 0.29 | 6.91 | 0.66 | 8.31 | 1.72
| Analytical results for Ta in OREAS 284 (Certified Value 1.09 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.5855 | 1.15 | 0.2 | 5 | 1.82 | 1.1 | 1.05 | 1.15 | 1.04 | 1.11 | 1.14 | 1.0236 | 1.5 | 1.1 | 1.02 | 1.09 | 1.1 | 0.98 | 1.13 | 1.12 | < 1 | 0.92 | 1.14 |  |  |  | 1
| 1 | 2 | 1.4546 | 1.14 | < 0.1 | 5 | 1.84 | 1 | 1.06 | 1.1 | 1.02 | 1.11 | 1.17 | 1.1935 | 1.3 | 1.2 | 1.02 | 1.07 | 1 | 1.1 | 1.09 | 1.1 | < 1 | 0.96 | 1.16 |  |  |  | 25
| 1 | 3 | 1.373 | 1.12 | < 0.1 | 5 | 1.82 | 1 | 1.04 | 1.06 | 1.02 | 1.24 | 1.12 | 1.0543 | 1.3 | 1.17 | 1.03 | 1.07 | 1 | 1.03 | 1.11 | 1.11 | < 1 | 0.739 | 1.1 |  |  |  | 16
| 1 | 4 | 1.3713 | 1.16 | < 0.1 | 5 | 1.86 | 1 | 1.04 | 1.18 | 1.05 | 1.22 | 1.19 | 1.1278 | 1.2 | 1.1 | 1.03 | 1.1 | 1 | 1.09 | 1.1 | 1.12 | < 1 | 0.94 | 1.14 |  |  |  | 1.09
| 1 | 5 | 1.7336 | 1.12 | 1 | 5 | 1.88 | 1 | 1.11 | 1.1 | 1.04 | 1.13 | 1.14 | 1.1298 | 1.3 | 1.05 | 1.02 | 1.07 | 1 | 1.11 | 1.12 | 1.11 | < 1 | 0.86 | 1.11 |  |  |  | 58
| 1 | 6 | 1.3612 | 1.13 | < 0.1 | 5 | 1.99 | 0.9 | 1.05 | 1.14 | 1.02 | 1.11 | 1.16 | 1.1457 | 1.2 | 1.14 | 1.02 | 1.1 | 1 | 1.07 | 1.1 | 1.14 | < 1 | 0.97 | 1.1
| Mean |  | 1.4799 | 1.1367 | 0.6 | 5 | 1.8683 | 1 | 1.0583 | 1.1217 | 1.0317 | 1.1533 | 1.1533 | 1.1125 | 1.3 | 1.1267 | 1.0233 | 1.0833 | 1.0167 | 1.0633 | 1.1083 | 1.1167 |  | 0.89817 | 1.125
| Median |  | 1.4138 | 1.135 | 0.6 | 5 | 1.85 | 1 | 1.05 | 1.12 | 1.03 | 1.12 | 1.15 | 1.1288 | 1.3 | 1.12 | 1.02 | 1.08 | 1 | 1.08 | 1.105 | 1.115 |  | 0.93 | 1.125
| Std Dev. |  | 0.15065 | 0.01633 | 0.56569 | 0 | 0.064005 | 0.063246 | 0.026394 | 0.043089 | 0.013292 | 0.060222 | 0.025033 | 0.062434 | 0.10954 | 0.054283 | 0.005164 | 0.015055 | 0.040825 | 0.049666 | 0.01472 | 0.013663 |  | 0.087179 | 0.0251
| Rel.Std.Dev. |  | 10.18% | 1.437% | 94.28% | 0% | 3.426% | 6.325% | 2.494% | 3.842% | 1.288% | 5.222% | 2.171% | 5.612% | 8.427% | 4.818% | 0.5046% | 1.39% | 4.016% | 4.671% | 1.328% | 1.224% |  | 9.706% | 2.231%
| PDM3 |  | 35.77% | 4.279% | -44.96% | 358.7% | 71.4% | -8.259% | -2.908% | 2.903% | -5.354% | 5.808% | 5.808% | 2.058% | 19.26% | 3.361% | -6.119% | -0.6142% | -6.73% | -2.449% | 1.679% | 2.444% |  | -17.6% | 3.208%
| Z-Score (Absolute) |  | 5.03 | 0.36 | 11.93 | NA | 10.31 | 1.5 | 0.71 | 0.15 | 1.07 | 0.58 | 0.58 | 0.03 | 2.58 | 0.22 | 1.18 | 0.37 | 1.28 | 0.64 | 0.03 | 0.09 | 8.3 | 2.89 | 0.2
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 284 (Certified Value 0.79 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab; 1.17: Lab |  |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.11: 1.10; 1.12: 1.13; 1.13: 1.21; 1.14: 1.25; 1.15: 1.26; 1.16: 1.29; 1.17: 1.32 |  |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*OES/MS; 1.14: 4A*MS; 1.15: 4A*MS; 1.16: 4A*OES/MS; 1.17: 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.03: 0.82266; 1.04: 0.78; 1.09: 0.8; 1.11: 0.7759; 1.12: 0.8; 1.13: 0.8; 1.14: 0.75; 1.15: 0.86; 1.16: < 1; 1.17: 0.69 |  |  |  | 1
| 1 | 2 | 1.03: 0.82911; 1.04: 0.79; 1.09: 0.8; 1.11: 0.7714; 1.12: 0.9; 1.13: 0.88; 1.14: 0.8; 1.15: 0.84; 1.16: < 1; 1.17: 0.67 |  |  |  | 26
| 1 | 3 | 1.03: 0.84884; 1.04: 0.8; 1.09: 0.9; 1.11: 0.7689; 1.12: 0.8; 1.13: 0.86; 1.14: 0.76; 1.15: 0.84; 1.16: < 1; 1.17: 0.65 |  |  |  | 16
| 1 | 4 | 1.03: 0.87366; 1.04: 0.79; 1.09: 0.8; 1.11: 0.7197; 1.12: 0.8; 1.13: 0.84; 1.14: 0.8; 1.15: 0.86; 1.16: < 1; 1.17: 0.67 |  |  |  | 0.79188
| 1 | 5 | 1.03: 0.8431; 1.04: 0.8; 1.09: 0.8; 1.11: 0.7334; 1.12: 0.8; 1.13: 0.83; 1.14: 0.82; 1.15: 0.9; 1.16: < 1; 1.17: 0.66 |  |  |  | 59
| 1 | 6 | 1.03: 0.83827; 1.04: 0.81; 1.09: 0.8; 1.11: 0.7141; 1.12: 0.7; 1.13: 0.83; 1.14: 0.79; 1.15: 0.86; 1.16: < 1; 1.17: 0.69
| Mean |  | 1.03: 0.84261; 1.04: 0.795; 1.09: 0.81667; 1.11: 0.74723; 1.12: 0.8; 1.13: 0.84; 1.14: 0.78667; 1.15: 0.86; 1.17: 0.67167
| Median |  | 1.03: 0.84068; 1.04: 0.795; 1.09: 0.8; 1.11: 0.75115; 1.12: 0.8; 1.13: 0.835; 1.14: 0.795; 1.15: 0.86; 1.17: 0.67
| Std Dev. |  | 1.03: 0.017902; 1.04: 0.010488; 1.09: 0.040825; 1.11: 0.028009; 1.12: 0.063246; 1.13: 0.027568; 1.14: 0.026583; 1.15: 0.021909; 1.17: 0.016021
| Rel.Std.Dev. |  | 1.03: 2.125%; 1.04: 1.319%; 1.09: 4.999%; 1.11: 3.748%; 1.12: 7.906%; 1.13: 3.282%; 1.14: 3.379%; 1.15: 2.548%; 1.17: 2.385%
| PDM3 |  | 1.03: 6.406%; 1.04: 0.3937%; 1.09: 3.13%; 1.11: -5.638%; 1.12: 1.025%; 1.13: 6.076%; 1.14: -0.6586%; 1.15: 8.602%; 1.17: -15.18%
| Z-Score (Absolute) |  | 1.03: 0.69; 1.04: 0.06; 1.09: NA; 1.11: 0.58; 1.12: NA; 1.13: 0.66; 1.14: 0.06; 1.15: 0.92; 1.16: 3.89; 1.17: 1.59
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 284 (Certified Value 0.056 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | 0.1 | 5 | 0.09 | 0.09 | < 0.05 | < 0.05 | < 0.05 | 0.07 | 0.06 | < 0.1 | < 0.1 | 0.08 | < 0.1 | 0.03 | < 0.5 | < 0.05 | < 0.2 | 0.18 | < 0.05 | [5 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.1 | 5 | 0.1 | 0.07 | 0.1 | < 0.05 | 0.06 | 0.06 | 0.18 | < 0.1 | < 0.1 | 0.07 | < 0.1 | 0.08 | < 0.5 | 0.05 | < 0.2 | 0.24 | < 0.05 | [5 empty] | 27
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 5 | 0.06 | 0.15 | 0.05 | 0.05 | 0.05 | 0.08 | < 0.05 | < 0.1 | < 0.1 | 0.08 | < 0.1 | 0.1 | < 0.5 | 0.07 | < 0.2 | < 0.01 | < 0.05 | [5 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 5 | 0.04 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.08 | < 0.05 | < 0.1 | < 0.1 | 0.05 | < 0.1 | 0.07 | < 0.5 | 0.18 | < 0.2 | < 0.01 | 0.12 | [5 empty] | 0.056015
| 1 | 5 | < 0.2 | < 0.2 | 0.2 | 5 | 0.07 | 0.05 | 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | 0.2 | 0.06 | < 0.1 | 0.1 | < 0.5 | 0.06 | < 0.2 | 0.19 | 0.05 | [5 empty] | 60
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 5 | 0.07 | 0.08 | 0.06 | 0.07 | 0.06 | 0.08 | < 0.05 | < 0.1 | < 0.1 | 0.05 | < 0.1 | 0.05 | < 0.5 | 0.06 | < 0.2 | 0.1 | 0.07
| Mean |  |  |  | 0.15 | 5 | 0.071667 | 0.088 | 0.065 | 0.056667 | 0.056667 | 0.074 | 0.12 |  | 0.2 | 0.065 |  | 0.071667 |  | 0.084 |  | 0.1775 | 0.08
| Median |  |  |  | 0.15 | 5 | 0.07 | 0.08 | 0.055 | 0.05 | 0.06 | 0.08 | 0.12 |  | 0.2 | 0.065 |  | 0.075 |  | 0.06 |  | 0.185 | 0.07
| Std Dev. |  |  |  | 0.070711 | 0 | 0.02137 | 0.037683 | 0.023805 | 0.011547 | 0.0057735 | 0.0089443 | 0.084853 |  |  | 0.013784 |  | 0.027869 |  | 0.054129 |  | 0.057951 | 0.036056
| Rel.Std.Dev. |  |  |  | 47.14% | 0% | 29.82% | 42.82% | 36.62% | 20.38% | 10.19% | 12.09% | 70.71% |  |  | 21.21% |  | 38.89% |  | 64.44% |  | 32.65% | 45.07%
| PDM3 |  |  |  | 167.8% | 8826% | 27.94% | 57.1% | 16.04% | 1.163% | 1.163% | 32.11% | 114.2% |  | 257% | 16.04% |  | 27.94% |  | 49.96% |  | 216.9% | 42.82%
| Z-Score (Absolute) |  | 1.18 | 1.18 | NA | NA | 0.11 | 0.33 | 0.64 | 1.05 | 1.05 | 0.11 | 0.45 | 0.71 | NA | 0.14 | 0.71 | 0.11 | 6.82 | 0.2 | 1.18 | 1.93 | 0.61
| Analytical results for Th in OREAS 284 (Certified Value 15.3 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 15.5423 | 15.21 | 15.7 | 34.9972 | 10.59 | 13.9 | 14.65 | 13.8 | 15.95 | 14.2 | 15.45 | 16.4267 | 17.9 | 13.8 | 14.8 | 15.5 | 15.3 | 14.27 | 16.05 | 15.21 | 14 | 17.6 | 12 |  |  |  | 1
| 1 | 2 | 14.9864 | 15.78 | 15.9 | 14.5032 | 10.7 | 16.8 | 15.15 | 15 | 14.2 | 14.2 | 15.35 | 16.1825 | 18.3 | 13.9 | 15 | 15.6 | 15.6 | 15.52 | 15.79 | 15.3 | 14 | 17 | 10.6 |  |  |  | 28
| 1 | 3 | 15.529 | 15.58 | 16.3 | 15.2531 | 10.84 | 15.4 | 14.6 | 14.9 | 15.5 | 15 | 14.85 | 15.9895 | 17.7 | 13.5 | 15.2 | 16.1 | 15 | 15.33 | 15.81 | 15.37 | 14 | 16.7 | 10.2 |  |  |  | 16
| 1 | 4 | 16.331 | 15.37 | 15.2 | 14.1185 | 10.99 | 15.2 | 14.65 | 15.8 | 14.9 | 14.6 | 15.9 | 16.2602 | 17.9 | 14 | 15.2 | 15.8 | 14.8 | 15.22 | 16.23 | 15.59 | 13 | 17.6 | 10.4 |  |  |  | 15.2733
| 1 | 5 | 15.9731 | 15.71 | 14.3 | 13.469 | 10.93 | 14.2 | 16.4 | 15.15 | 15.2 | 14.15 | 15.2 | 15.9415 | 16.1 | 13.4 | 15 | 15.9 | 15.4 | 15.75 | 15.89 | 15.29 | 13 | 17 | 10.6 |  |  |  | 61
| 1 | 6 | 15.2067 | 15.68 | 16 | 14.8184 | 10.75 | 14 | 15 | 16.55 | 13.9 | 14.3 | 14.9 | 16.2952 | 16.6 | 13.3 | 15.2 | 15.6 | 14.9 | 15.83 | 16.45 | 15.23 | 14 | 18.3 | 10.7
| Mean |  | 15.5947 | 15.555 | 15.5667 | 17.8599 | 10.8 | 14.9167 | 15.075 | 15.2 | 14.9417 | 14.4083 | 15.275 | 16.1826 | 17.4167 | 13.65 | 15.0667 | 15.75 | 15.1667 | 15.32 | 16.0367 | 15.3317 | 13.6667 | 17.3667 | 10.75
| Median |  | 15.5357 | 15.63 | 15.8 | 14.6608 | 10.795 | 14.7 | 14.825 | 15.075 | 15.05 | 14.25 | 15.275 | 16.2214 | 17.8 | 13.65 | 15.1 | 15.7 | 15.15 | 15.425 | 15.97 | 15.295 | 14 | 17.3 | 10.6
| Std Dev. |  | 0.49253 | 0.22098 | 0.72019 | 8.4176 | 0.14913 | 1.1179 | 0.68611 | 0.92466 | 0.77873 | 0.33229 | 0.38827 | 0.18639 | 0.86352 | 0.2881 | 0.1633 | 0.22583 | 0.31411 | 0.56526 | 0.26158 | 0.13862 | 0.5164 | 0.58195 | 0.63797
| Rel.Std.Dev. |  | 3.158% | 1.421% | 4.626% | 47.13% | 1.381% | 7.494% | 4.551% | 6.083% | 5.212% | 2.306% | 2.542% | 1.152% | 4.958% | 2.111% | 1.084% | 1.434% | 2.071% | 3.69% | 1.631% | 0.9042% | 3.779% | 3.351% | 5.935%
| PDM3 |  | 2.104% | 1.844% | 1.921% | 16.94% | -29.29% | -2.335% | -1.298% | -0.4801% | -2.171% | -5.663% | 0.011% | 5.953% | 14.03% | -10.63% | -1.353% | 3.121% | -0.6983% | 0.3056% | 4.998% | 0.382% | -10.52% | 13.71% | -29.62%
| Z-Score (Absolute) |  | 0.54 | 0.47 | 0.49 | 4.69 | 8.23 | 0.7 | 0.41 | 0.18 | 0.65 | 1.63 | 0.04 | 1.62 | 3.88 | 3.02 | 0.42 | 0.83 | 0.24 | 0.04 | 1.35 | 0.06 | NA | 3.79 | 8.33
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 284 (Certified Value 0.373 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.36944 | 0.36932 | 0.3714 | 0.347 | 0.335 | 0.41 | 0.3753 | 0.374 | 0.36 | 0.389 | 0.357 | 0.354 | 0.35 | 0.38591 | 0.386 | 0.36 | 0.372 | 0.36 | 0.417 | 0.36 | 0.3683 | 0.3714 | 0.3593 | 0.396 | 0.372 | 0.36 | 1
| 1 | 2 | 0.36851 | 0.37784 | 0.3688 | 0.325 | 0.3373 | 0.41 | 0.3767 | 0.368 | 0.371 | 0.378 | 0.354 | 0.367 | 0.371 | 0.38803 | 0.391 | 0.37 | 0.384 | 0.36 | 0.38 | 0.381 | 0.366 | 0.3789 |  | 0.402 | 0.379 | 0.36 | 29
| 1 | 3 | 0.36794 | 0.38092 | 0.3796 | 0.322 | 0.3405 | 0.41 | 0.3755 | 0.378 | 0.359 | 0.374 | 0.351 | 0.372 | 0.363 | 0.38032 | 0.384 | 0.37 | 0.372 | 0.35 | 0.401 | 0.361 | 0.3647 | 0.3781 | 0.35203 | 0.39 | 0.375 | 0.37 | 16
| 1 | 4 | 0.36226 | 0.38199 | 0.3731 | 0.24 | 0.3435 | 0.41 | 0.3832 | 0.357 | 0.366 | 0.375 | 0.362 | 0.375 | 0.361 | 0.38095 | 0.384 | 0.37 | 0.382 | 0.36 | 0.389 | 0.379 | 0.3689 | 0.3794 | 0.36465 | 0.384 | 0.392 | 0.37 | 0.37309
| 1 | 5 | 0.36211 | 0.37474 | 0.3698 | 0.376 | 0.3336 | 0.41 | 0.374 | 0.373 | 0.372 | 0.383 | 0.356 | 0.372 | 0.367 | 0.37861 | 0.363 | 0.38 | 0.37 | 0.35 | 0.395 | 0.383 | 0.3709 | 0.3826 | 0.35611 | 0.396 | 0.387 | 0.37 | 62
| 1 | 6 | 0.36358 | 0.37666 | 0.3791 | 0.253 | 0.3392 | 0.42 | 0.3747 | 0.356 | 0.367 | 0.382 | 0.354 | 0.364 | 0.372 | 0.38389 | 0.385 | 0.37 | 0.375 | 0.35 | 0.387 | 0.385 | 0.3655 | 0.3802 | 0.36085 | 0.408 | 0.401 | 0.38
| Mean |  | 0.36564 | 0.37691 | 0.37363 | 0.3105 | 0.33818 | 0.41167 | 0.37657 | 0.36767 | 0.36583 | 0.38017 | 0.35567 | 0.36733 | 0.364 | 0.38295 | 0.38217 | 0.37 | 0.37583 | 0.355 | 0.39483 | 0.37483 | 0.36738 | 0.37843 | 0.35859 | 0.396 | 0.38433 | 0.36833
| Median |  | 0.36576 | 0.37725 | 0.37225 | 0.3235 | 0.33825 | 0.41 | 0.3754 | 0.3705 | 0.3665 | 0.38 | 0.355 | 0.3695 | 0.365 | 0.38242 | 0.3845 | 0.37 | 0.3735 | 0.355 | 0.392 | 0.38 | 0.36715 | 0.37915 | 0.3593 | 0.396 | 0.383 | 0.37
| Std Dev. |  | 0.0033503 | 0.0045846 | 0.004665 | 0.053369 | 0.0036493 | 0.0040825 | 0.0033714 | 0.0092232 | 0.0054191 | 0.0056362 | 0.0037238 | 0.0076333 | 0.0080994 | 0.0036095 | 0.0097451 | 0.0063246 | 0.0058109 | 0.0054772 | 0.012999 | 0.011286 | 0.0023735 | 0.0037739 | 0.004784 | 0.0084853 | 0.011057 | 0.0075277
| Rel.Std.Dev. |  | 0.9163% | 1.216% | 1.249% | 17.19% | 1.079% | 0.9917% | 0.8953% | 2.509% | 1.481% | 1.483% | 1.047% | 2.078% | 2.225% | 0.9425% | 2.55% | 1.709% | 1.546% | 1.543% | 3.292% | 3.011% | 0.6461% | 0.9973% | 1.334% | 2.143% | 2.877% | 2.044%
| PDM3 |  | -1.998% | 1.023% | 0.1445% | -16.78% | -9.357% | 10.34% | 0.9307% | -1.455% | -1.946% | 1.896% | -4.671% | -1.544% | -2.437% | 2.643% | 2.432% | -0.8293% | 0.7342% | -4.85% | 5.827% | 0.4661% | -1.531% | 1.431% | -3.888% | 6.139% | 3.012% | -1.276%
| Z-Score (Absolute) |  | 0.58 | 0.48 | 0.17 | 5.73 | 3.14 | 3.72 | 0.44 | 0.39 | 0.56 | 0.78 | 1.51 | 0.42 | 0.73 | 1.04 | 0.97 | 0.17 | 0.38 | 1.57 | 2.15 | 0.28 | 0.41 | 0.62 | 1.24 | 2.23 | 1.17 | 0.33
| Analytical results for Tl in OREAS 284 (Certified Value 0.95 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.95346 | 0.98 | 0.92 | 1.028 | 0.88 | 0.93 | 0.99 | 0.9 | 0.86 | 0.87 | 0.96614 | 0.7 | 0.93 | 0.9 | 0.99 | 1 | 0.91 | 0.99 | 0.95 | 0.99 | 0.93 | [5 empty] | 1
| 1 | 2 | 0.92969 | 1.03 | 0.92 | 1.028 | 0.98 | 0.92 | 0.98 | 0.91 | 0.83 | 0.9 | 1.0949 | 0.8 | 0.96 | 0.9 | 0.96 | 1 | 1 | 0.97 | 0.98 | 0.95 | 0.93 | [5 empty] | 30
| 1 | 3 | 0.97447 | 1.02 | 0.95 | 1.052 | 0.94 | 0.93 | 0.96 | 0.89 | 0.86 | 0.89 | 0.9306 | 0.8 | 0.94 | 0.9 | 1.01 | 1 | 0.94 | 0.96 | 0.97 | 0.96 | 0.92 | [5 empty] | 16
| 1 | 4 | 0.96316 | 0.97 | 0.95 | 1.076 | 0.93 | 0.91 | 1.01 | 0.94 | 0.88 | 0.94 | 1.0704 | 0.8 | 0.93 | 0.9 | 1.02 | 1 | 0.97 | 0.97 | 0.94 | 0.98 | 0.94 | [5 empty] | 0.95386
| 1 | 5 | 0.9564 | 1 | 0.93 | 1.042 | 0.89 | 0.96 | 0.99 | 0.93 | 0.85 | 0.92 | 1.0217 | 0.7 | 0.94 | 0.91 | 0.97 | 1 | 1 | 0.98 | 0.93 | 0.95 | 0.94 | [5 empty] | 63
| 1 | 6 | 0.94636 | 0.99 | 0.96 | 1.064 | 0.9 | 0.94 | 1.04 | 0.91 | 0.87 | 0.91 | 1.0032 | 0.8 | 0.95 | 0.91 | 0.99 | 1 | 1 | 0.99 | 0.98 | 1.01 | 0.94
| Mean |  | 0.95392 | 0.99833 | 0.93833 | 1.0483 | 0.92 | 0.93167 | 0.995 | 0.91333 | 0.85833 | 0.905 | 1.0145 | 0.76667 | 0.94167 | 0.90333 | 0.99 | 1 | 0.97 | 0.97667 | 0.95833 | 0.97333 | 0.93333
| Median |  | 0.95493 | 0.995 | 0.94 | 1.047 | 0.915 | 0.93 | 0.99 | 0.91 | 0.86 | 0.905 | 1.0124 | 0.8 | 0.94 | 0.9 | 0.99 | 1 | 0.985 | 0.975 | 0.96 | 0.97 | 0.935
| Std Dev. |  | 0.015224 | 0.023166 | 0.017224 | 0.019449 | 0.037417 | 0.017224 | 0.027386 | 0.018619 | 0.017224 | 0.02429 | 0.061898 | 0.05164 | 0.01169 | 0.005164 | 0.022804 | 0 | 0.037947 | 0.012111 | 0.02137 | 0.024221 | 0.008165
| Rel.Std.Dev. |  | 1.596% | 2.32% | 1.836% | 1.855% | 4.067% | 1.849% | 2.752% | 2.039% | 2.007% | 2.684% | 6.101% | 6.736% | 1.241% | 0.5717% | 2.303% | 0% | 3.912% | 1.24% | 2.23% | 2.488% | 0.8748%
| PDM3 |  | 0.006153% | 4.662% | -1.628% | 9.904% | -3.55% | -2.327% | 4.313% | -4.249% | -10.02% | -5.123% | 6.357% | -19.63% | -1.279% | -5.297% | 3.788% | 4.837% | 1.692% | 2.391% | 0.4685% | 2.041% | -2.152%
| Z-Score (Absolute) |  | 0.13 | 1.08 | 0.2 | 2.14 | 0.59 | 0.34 | 1.01 | 0.73 | 1.9 | 0.91 | 1.42 | 3.86 | 0.13 | 0.95 | 0.9 | NA | 0.47 | 0.61 | 0.22 | 0.54 | 0.31
Results from laboratory 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 284 (Certified Value 0.19 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab | [5 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.12: 1.10; 1.13: 1.13; 1.14: 1.26 | [5 empty] | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS; 1.14: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 1.03: 0.17654; 1.04: 0.19; 1.09: 0.2; 1.12: 0.1683; 1.13: 0.2; 1.14: 0.19 | [5 empty] | 1
| 1 | 2 | 1.03: 0.17763; 1.04: 0.19; 1.09: 0.2; 1.12: 0.1716; 1.13: 0.3; 1.14: 0.19 | [5 empty] | 5
| 1 | 3 | 1.03: 0.18344; 1.04: 0.21; 1.09: 0.2; 1.12: 0.1706; 1.13: 0.2; 1.14: 0.2 | [5 empty] | 16
| 1 | 4 | 1.03: 0.18805; 1.04: 0.18; 1.09: 0.2; 1.12: 0.1619; 1.13: 0.2; 1.14: 0.2 | [5 empty] | 0.18907
| 1 | 5 | 1.03: 0.17488; 1.04: 0.18; 1.09: 0.2; 1.12: 0.1577; 1.13: 0.2; 1.14: 0.19 | [5 empty] | 64
| 1 | 6 | 1.03: 0.18157; 1.04: 0.2; 1.09: 0.2; 1.12: 0.1642; 1.13: 0.2; 1.14: 0.21
| Mean |  | 1.03: 0.18035; 1.04: 0.19167; 1.09: 0.2; 1.12: 0.16572; 1.13: 0.21667; 1.14: 0.19667
| Median |  | 1.03: 0.1796; 1.04: 0.19; 1.09: 0.2; 1.12: 0.16625; 1.13: 0.2; 1.14: 0.195
| Std Dev. |  | 1.03: 0.0049414; 1.04: 0.01169; 1.09: 3.0405e-17; 1.12: 0.0054087; 1.13: 0.040825; 1.14: 0.008165
| Rel.Std.Dev. |  | 1.03: 2.74%; 1.04: 6.099%; 1.09: 1.52e-14%; 1.12: 3.264%; 1.13: 18.84%; 1.14: 4.152%
| PDM3 |  | 1.03: -4.61%; 1.04: 1.375%; 1.09: 5.783%; 1.12: -12.35%; 1.13: 14.6%; 1.14: 4.02%
| Z-Score (Absolute) |  | 1.03: 0.95; 1.04: 0.17; 1.09: 0.4; 1.12: 1.95; 1.13: 1.54; 1.14: 0.17
| Analytical results for U in OREAS 284 (Certified Value 4.09 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 3.6497 | 3.81 | 3.9 | 3.8823 | 4.31 | 4 | 3.9 | 4.3 | 3.6 | 3.8 | 3.9 | 4.2418 | 4.6 | 4.1 | 4 | 3.76 | 4.6 | 4.66 | 4.62 | 4.43 | 2 | 4.36 | 3.1 |  |  |  | 1
| 1 | 2 | 3.6972 | 4.65 | 4.2 | 3.8452 | 4.19 | 4.5 | 4 | 4.1 | 3.2 | 3.8 | 4.2 | 4.2868 | 4.9 | 4 | 3.74 | 3.66 | 4.1 | 4.54 | 4.71 | 4.23 | < 2 | 4.56 | 3.4 |  |  |  | 32
| 1 | 3 | 3.8654 | 4.54 | 4.6 | 3.8618 | 4.3 | 3.9 | 4 | 4.4 | 3.6 | 5.1 | 3.7 | 4.2185 | 5.4 | 3.9 | 3.88 | 4.08 | 3.5 | 4.15 | 4.72 | 3.89 | < 2 | 4.25 | 3.3 |  |  |  | 16
| 1 | 4 | 3.9344 | 4.58 | 3.9 | 3.2573 | 4.73 | 4 | 4.1 | 4.6 | 3.9 | 3.5 | 4.5 | 4.0027 | 4.6 | 3.9 | 4 | 3.83 | 3.7 | 4.05 | 4.62 | 4.16 | < 2 | 4.71 | 3.5 |  |  |  | 4.0929
| 1 | 5 | 3.7168 | 4.18 | 3.2 | 3.735 | 4.29 | 3.6 | 4.2 | 5 | 3.5 | 3.4 | 4.7 | 4.0347 | 4.2 | 4 | 3.75 | 4.18 | 4.4 | 3.76 | 4.62 | 4.16 | < 2 | 3.85 | 2.9 |  |  |  | 65
| 1 | 6 | 3.8295 | 4.56 | 4.4 | 3.9067 | 4.74 | 4.5 | 4.7 | 4.6 | 4.1 | 3.8 | 4 | 4.0137 | 4.8 | 4.1 | 3.64 | 4.41 | 3.5 | 4.11 | 4.01 | 3.83 | < 2 | 4.95 | 3.3
| Mean |  | 3.7822 | 4.3867 | 4.0333 | 3.7481 | 4.4267 | 4.0833 | 4.15 | 4.5 | 3.65 | 3.9 | 4.1667 | 4.133 | 4.75 | 4 | 3.835 | 3.9867 | 3.9667 | 4.2117 | 4.55 | 4.1167 | 2 | 4.4467 | 3.25
| Median |  | 3.7732 | 4.55 | 4.05 | 3.8535 | 4.305 | 4 | 4.05 | 4.5 | 3.6 | 3.8 | 4.1 | 4.1266 | 4.7 | 4 | 3.815 | 3.955 | 3.9 | 4.13 | 4.62 | 4.16 | 2 | 4.46 | 3.3
| Std Dev. |  | 0.11079 | 0.32727 | 0.49261 | 0.24762 | 0.24271 | 0.35449 | 0.2881 | 0.30984 | 0.31464 | 0.61319 | 0.37771 | 0.12937 | 0.39875 | 0.089443 | 0.14883 | 0.28563 | 0.47188 | 0.33259 | 0.26863 | 0.22286 |  | 0.38412 | 0.21679
| Rel.Std.Dev. |  | 2.929% | 7.461% | 12.21% | 6.607% | 5.483% | 8.682% | 6.942% | 6.885% | 8.62% | 15.72% | 9.065% | 3.13% | 8.395% | 2.236% | 3.881% | 7.165% | 11.9% | 7.897% | 5.904% | 5.414% |  | 8.638% | 6.671%
| PDM3 |  | -7.592% | 7.176% | -1.456% | -8.427% | 8.154% | -0.2348% | 1.394% | 9.945% | -10.82% | -4.714% | 1.801% | 0.9798% | 16.05% | -2.271% | -6.302% | -2.597% | -3.085% | 2.901% | 11.17% | 0.5796% | -51.14% | 8.642% | -20.6%
| Z-Score (Absolute) |  | 0.92 | 0.83 | 0.19 | 1.02 | 0.95 | 0.05 | 0.15 | 1.16 | 1.31 | 0.58 | 0.19 | 0.1 | 1.89 | 0.29 | 0.77 | 0.33 | 0.39 | 0.32 | 1.31 | 0.05 | NA | 1.01 | 2.47
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 284 (Certified Value 62 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
1
| 1 | 1 | 65.0628 | 63.0161 | 64 | 61 | 56.4475 | 62 | 62 | 67 | 62 | 64 | 60 | 59 | 60 | 62.5335 | 64 | 59 | 61 | 61.3 | 62 | 59 | 66 | 58 | 64.74 | 64.9 | 59 | 63 | 1
| 1 | 2 | 66.3462 | 63.8927 | 64 | 61 | 56.1564 | 63 | 62 | 64 | 62 | 61 | 59 | 61 | 65 | 64.4044 | 65 | 59 | 61 | 59.1 | 63 | 63 | 66 | 61 |  | 67.7 | 60 | 61 | 16
| 1 | 3 | 65.89 | 65.0766 | 65 | 60 | 56.0934 | 65 | 61 | 68 | 61 | 60 | 59 | 62 | 64 | 63.1246 | 65 | 60 | 60 | 58.8 | 60 | 61 | 66 | 63 | 64.42 | 63.4 | 59.3 | 61 | 16
| 1 | 4 | 65.6952 | 65.2736 | 65 | 44 | 57.114 | 63 | 63 | 64 | 61 | 63 | 60 | 63 | 63 | 65.1794 | 63 | 59 | 61 | 62.1 | 58 | 63 | 63 | 63 | 65.39 | 64.6 | 62 | 60 | 62.4928
| 1 | 5 | 66.6066 | 62.772 | 63 | 64 | 55.2367 | 63 | 61 | 67 | 63 | 64 | 60 | 62 | 64 | 63.2533 | 60 | 60 | 61 | 62.1 | 60 | 63 | 67 | 62 | 63.3 | 66.8 | 60 | 60 | 66
| 1 | 6 | 65.8998 | 63.7863 | 66 | 47 | 56.3496 | 64 | 62 | 65 | 62 | 63 | 59 | 62 | 64 | 64.996 | 65 | 59 | 61 | 62.8 | 58 | 64 | 66 | 61 | 63.62 | 68.6 | 63 | 61
| Mean |  | 65.9168 | 63.9696 | 64.5 | 56.1667 | 56.2329 | 63.3333 | 61.8333 | 65.8333 | 61.8333 | 62.5 | 59.5 | 61.5 | 63.3333 | 63.9152 | 63.6667 | 59.3333 | 60.8333 | 61.0333 | 60.1667 | 62.1667 | 65.6667 | 61.3333 | 64.294 | 66 | 60.55 | 61
| Median |  | 65.8949 | 63.8395 | 64.5 | 60.5 | 56.253 | 63 | 62 | 66 | 62 | 63 | 59.5 | 62 | 64 | 63.8289 | 64.5 | 59 | 61 | 61.7 | 60 | 63 | 66 | 61.5 | 64.42 | 65.85 | 60 | 61
| Std Dev. |  | 0.53659 | 1.0304 | 1.0488 | 8.4242 | 0.60889 | 1.0328 | 0.75277 | 1.7224 | 0.75277 | 1.6432 | 0.54772 | 1.3784 | 1.7512 | 1.0934 | 1.9664 | 0.5164 | 0.40825 | 1.6848 | 2.0412 | 1.8348 | 1.3663 | 1.8619 | 0.84533 | 2.011 | 1.5922 | 1.0954
| Rel.Std.Dev. |  | 0.814% | 1.611% | 1.626% | 15% | 1.083% | 1.631% | 1.217% | 2.616% | 1.217% | 2.629% | 0.9205% | 2.241% | 2.765% | 1.711% | 3.089% | 0.8703% | 0.6711% | 2.761% | 3.393% | 2.951% | 2.081% | 3.036% | 1.315% | 3.047% | 2.63% | 1.796%
| PDM3 |  | 5.479% | 2.363% | 3.212% | -10.12% | -10.02% | 1.345% | -1.055% | 5.346% | -1.055% | 0.01158% | -4.789% | -1.589% | 1.345% | 2.276% | 1.878% | -5.056% | -2.655% | -2.335% | -3.722% | -0.5218% | 5.079% | -1.855% | 2.882% | 5.612% | -3.109% | -2.389%
| Z-Score (Absolute) |  | 1.51 | 0.76 | 0.96 | 2.25 | 2.22 | 0.51 | 0.06 | 1.48 | 0.06 | 0.19 | 0.96 | 0.19 | 0.51 | 0.74 | 0.64 | 1.03 | 0.45 | 0.37 | 0.71 | 0.06 | 1.41 | 0.26 | 0.88 | 1.54 | 0.56 | 0.39
| Analytical results for W in OREAS 284 (Certified Value 16.1 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 16.1675 | 16.5 | 11.2 | 12.4603 | 24 | 17.61 | 14.1 | 16.7 | 18.7 | 16.5 | 16.4 | 16.5 | 16.0665 | 11.4 | 14.5 | 14.4 | 16.5 | 16.8 | 14.4 | 15.7 | 16.2 | 12 | 16.8 | 16.3 |  |  | 1
| 1 | 2 | 15.9288 | 16.4 | 6.4 | 13.8088 | 25 | 16.74 | 15 | 14.9 | 16 | 15.4 | 17.7 | 16.9 | 16.8462 | 11.4 | 14.7 | 14.7 | 16 | 16.5 | 16.8 | 16.8 | 16.7 | 11 | 16.5 | 16.2 |  |  | 34
| 1 | 3 | 16.9041 | 15.9 | 4.5 | 12.0081 | 25 | 17.87 | 14.9 | 15.4 | 15.2 | 16 | 17.4 | 16 | 15.9363 | 12.8 | 14.8 | 14.3 | 16.5 | 16.4 | 15.1 | 16.1 | 16.7 | 7 | 16.7 | 15.6 |  |  | 16
| 1 | 4 | 16.0552 | 17.7 | 1 | 12.7382 | 24 | 17.34 | 26.6 | 14.7 | 16.1 | 15.8 | 17 | 17 | 15.9954 | 12.3 | 14.9 | 14.3 | 19.5 | 17.1 | 15 | 19.3 | 16.5 | 8 | 16.1 | 16.2 |  |  | 16.0541
| 1 | 5 | 16.0548 | 16.2 | 15.2 | 14.9956 | 25 | 17.51 | 15 | 16.3 | 14.6 | 16.5 | 15.8 | 16.3 | 15.4139 | 10.8 | 15.4 | 14.4 | 17.7 | 16.8 | 15.4 | 17.9 | 16.2 | 10 | 16.3 | 14.5 |  |  | 67
| 1 | 6 | 15.653 | 15.9 | 0.9 | 14.811 | 25 | 17.71 | 14.4 | 14.6 | 16.5 | 15.5 | 17.6 | 16.4 | 16.5789 | 12.4 | 14.7 | 14.6 | 16.2 | 16.5 | 16.4 | 16.9 | 17.1 | 11 | 17 | 15.1
| Mean |  | 16.1272 | 16.4333 | 6.5333 | 13.4703 | 24.6667 | 17.4633 | 16.6667 | 15.4333 | 16.1833 | 15.95 | 16.9833 | 16.5167 | 16.1395 | 11.85 | 14.8333 | 14.45 | 17.0667 | 16.6833 | 15.5167 | 17.1167 | 16.5667 | 9.8333 | 16.5667 | 15.65
| Median |  | 16.055 | 16.3 | 5.45 | 13.2735 | 25 | 17.56 | 14.95 | 15.15 | 16.05 | 15.9 | 17.2 | 16.45 | 16.0309 | 11.85 | 14.75 | 14.4 | 16.5 | 16.65 | 15.25 | 16.85 | 16.6 | 10.5 | 16.6 | 15.9
| Std Dev. |  | 0.41951 | 0.66833 | 5.7165 | 1.2596 | 0.5164 | 0.39707 | 4.88 | 0.88015 | 1.4106 | 0.47645 | 0.74944 | 0.37639 | 0.50724 | 0.76354 | 0.30768 | 0.16432 | 1.3307 | 0.26394 | 0.90866 | 1.3091 | 0.34448 | 1.9408 | 0.33267 | 0.7287
| Rel.Std.Dev. |  | 2.601% | 4.067% | 87.5% | 9.351% | 2.094% | 2.274% | 29.28% | 5.703% | 8.716% | 2.987% | 4.413% | 2.279% | 3.143% | 6.443% | 2.074% | 1.137% | 7.797% | 1.582% | 5.856% | 7.648% | 2.079% | 19.74% | 2.008% | 4.656%
| PDM3 |  | 0.4555% | 2.362% | -59.3% | -16.09% | 53.65% | 8.778% | 3.816% | -3.867% | 0.8049% | -0.6485% | 5.788% | 2.881% | 0.532% | -26.19% | -7.604% | -9.992% | 6.307% | 3.919% | -3.348% | 6.619% | 3.193% | -38.75% | 3.193% | -2.517%
| Z-Score (Absolute) |  | 0.06 | 0.25 | 9.76 | 2.74 | 8.58 | 1.29 | 0.49 | 0.76 | 0 | 0.24 | 0.81 | 0.34 | 0.04 | 4.38 | 1.37 | 1.75 | 0.89 | 0.51 | 0.67 | 0.94 | 0.39 | NA | 0.39 | 0.54
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 284 (Certified Value 15.8 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 15.1778 | 15.81 | 17.1 | 15.543 | 16 | 14.7 | 14.7 | 16.2 | 15.8 | 16.7 | 13.5 | 17.3 | 17.7412 | 14.5 | 15 | 15.4 | 16.9 | 15.4 | 15.3 | 16.4 | 16.95 | 18 | 17 | 13.8 |  |  | 1
| 1 | 2 | 15.389 | 16.72 | 16.8 | 16.6349 | 15 | 14.92 | 15.6 | 16.9 | 16.2 | 14 | 14.4 | 17.1 | 18.0253 | 14.9 | 15.2 | 15.7 | 16.7 | 15.3 | 16.5 | 15.9 | 15.6 | 18 | 15.7 | 13.1 |  |  | 35
| 1 | 3 | 15.6015 | 16.22 | 17.8 | 16.4468 | 15 | 15.13 | 15.5 | 16 | 15.9 | 15.4 | 14.3 | 17.3 | 16.386 | 14.3 | 15.1 | 15.9 | 16.7 | 14.9 | 15.4 | 16.06 | 15.88 | 17 | 15.8 | 12.9 |  |  | 16
| 1 | 4 | 15.5935 | 16.1 | 16.7 | 16.5161 | 15 | 15.55 | 14.6 | 16.1 | 17.1 | 15.9 | 14.3 | 17.9 | 17.1175 | 14.9 | 15.2 | 15.9 | 17.2 | 14.7 | 16.2 | 16.28 | 15.72 | 17 | 16.7 | 13.5 |  |  | 15.7883
| 1 | 5 | 15.2879 | 16.1 | 16.1 | 15.2132 | 15 | 15.42 | 15 | 17 | 16.3 | 15.5 | 13.6 | 17.2 | 16.7107 | 13.4 | 15.6 | 15.6 | 17.4 | 15.3 | 16.5 | 16.41 | 15.73 | 18 | 16.2 | 13.1 |  |  | 68
| 1 | 6 | 15.697 | 15.95 | 17.5 | 15.9085 | 15 | 15.4 | 14.2 | 16.3 | 17.7 | 14.7 | 14.2 | 18.2 | 17.3095 | 14.6 | 15.1 | 15.9 | 17.2 | 14.7 | 16.4 | 16.07 | 15.95 | 19 | 18.7 | 13.4
| Mean |  | 15.4578 | 16.15 | 17 | 16.0437 | 15.1667 | 15.1867 | 14.9333 | 16.4167 | 16.5 | 15.3667 | 14.05 | 17.5 | 17.215 | 14.4333 | 15.2 | 15.7333 | 17.0167 | 15.05 | 16.05 | 16.1867 | 15.9717 | 17.8333 | 16.6833 | 13.3
| Median |  | 15.4913 | 16.1 | 16.95 | 16.1776 | 15 | 15.265 | 14.85 | 16.25 | 16.25 | 15.45 | 14.25 | 17.3 | 17.2135 | 14.55 | 15.15 | 15.8 | 17.05 | 15.1 | 16.3 | 16.175 | 15.805 | 18 | 16.45 | 13.25
| Std Dev. |  | 0.2041 | 0.31318 | 0.60663 | 0.58204 | 0.40825 | 0.32959 | 0.54283 | 0.42622 | 0.74565 | 0.93737 | 0.3937 | 0.44272 | 0.61538 | 0.55737 | 0.20976 | 0.20656 | 0.29269 | 0.32094 | 0.55408 | 0.20781 | 0.49515 | 0.75277 | 1.1089 | 0.32863
| Rel.Std.Dev. |  | 1.32% | 1.939% | 3.568% | 3.628% | 2.692% | 2.17% | 3.635% | 2.596% | 4.519% | 6.1% | 2.802% | 2.53% | 3.575% | 3.862% | 1.38% | 1.313% | 1.72% | 2.132% | 3.452% | 1.284% | 3.1% | 4.221% | 6.647% | 2.471%
| PDM3 |  | -2.093% | 2.291% | 7.675% | 1.618% | -3.937% | -3.81% | -5.415% | 3.98% | 4.508% | -2.67% | -11.01% | 10.84% | 9.037% | -8.582% | -3.726% | -0.3479% | 7.78% | -4.676% | 1.658% | 2.523% | 1.162% | 12.95% | 5.669% | -15.76%
| Z-Score (Absolute) |  | 0.5 | 0.13 | 0.9 | 0.03 | NA | 0.75 | 0.98 | 0.37 | 0.45 | 0.58 | 1.78 | 1.36 | 1.1 | 1.43 | 0.73 | 0.25 | 0.92 | 0.87 | 0.04 | 0.16 | 0.03 | NA | 0.61 | 2.46
Results from laboratories 1.11 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 284 (Certified Value 1.17 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.13; 1.12: 1.21; 1.13: 1.25; 1.14: 1.26; 1.15: 1.29; 1.16: 1.32 |  |  | ppm
|  |  | 1.03: 4A*MS; 1.04: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*MS; 1.15: 4A*OES/MS; 1.16: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.03: 1.2037; 1.04: 1.12; 1.09: 1.4; 1.10: 1.085; 1.11: 0.9; 1.12: 1.1; 1.13: 1.1; 1.14: 1.22; 1.15: 1.4; 1.16: 1.1 |  |  | 1
| 1 | 2 | 1.03: 1.1712; 1.04: 1.16; 1.09: 1.4; 1.10: 1.0856; 1.11: 1.6; 1.12: 1.2; 1.13: 1.1; 1.14: 1.34; 1.15: 1.5; 1.16: 1 |  |  | 36
| 1 | 3 | 1.03: 1.2224; 1.04: 1.18; 1.09: 1.4; 1.10: 1.0975; 1.11: 1; 1.12: 1.2; 1.13: 1; 1.14: 1.19; 1.15: 1.4; 1.16: 1 |  |  | 16
| 1 | 4 | 1.03: 1.1655; 1.04: 1.13; 1.09: 1.3; 1.10: 1.1642; 1.11: 1.1; 1.12: 1.2; 1.13: 1.1; 1.14: 1.39; 1.15: 1.4; 1.16: 1.1 |  |  | 1.1739
| 1 | 5 | 1.03: 1.1667; 1.04: 1.1; 1.09: 1.3; 1.10: 0.9881; 1.11: 0.9; 1.12: 1.2; 1.13: 1.2; 1.14: 1.27; 1.15: 1.4; 1.16: 1 |  |  | 69
| 1 | 6 | 1.03: 1.1818; 1.04: 1.1; 1.09: 1.4; 1.10: 1.0226; 1.11: 0.9; 1.12: 1.2; 1.13: 1.1; 1.14: 1.22; 1.15: 1.4; 1.16: 1.1
| Mean |  | 1.03: 1.1852; 1.04: 1.1317; 1.09: 1.3667; 1.10: 1.0738; 1.11: 1.0667; 1.12: 1.1833; 1.13: 1.1; 1.14: 1.2717; 1.15: 1.4167; 1.16: 1.05
| Median |  | 1.03: 1.1765; 1.04: 1.125; 1.09: 1.4; 1.10: 1.0853; 1.11: 0.95; 1.12: 1.2; 1.13: 1.1; 1.14: 1.245; 1.15: 1.4; 1.16: 1.05
| Std Dev. |  | 1.03: 0.023102; 1.04: 0.032506; 1.09: 0.05164; 1.10: 0.061586; 1.11: 0.27325; 1.12: 0.040825; 1.13: 0.063246; 1.14: 0.078337; 1.15: 0.040825; 1.16: 0.054772
| Rel.Std.Dev. |  | 1.03: 1.949%; 1.04: 2.872%; 1.09: 3.779%; 1.10: 5.735%; 1.11: 25.62%; 1.12: 3.45%; 1.13: 5.75%; 1.14: 6.16%; 1.15: 2.882%; 1.16: 5.216%
| PDM3 |  | 1.03: 0.9633%; 1.04: -3.598%; 1.09: 16.42%; 1.10: -8.525%; 1.11: -9.135%; 1.12: 0.8032%; 1.13: -6.296%; 1.14: 8.328%; 1.15: 20.68%; 1.16: -10.55%
| Z-Score (Absolute) |  | 1.03: 0.21; 1.04: 0.2; 1.09: 1.62; 1.10: 0.65; 1.11: 0.7; 1.12: 0.2; 1.13: 0.44; 1.14: 0.88; 1.15: 2; 1.16: 0.83
| Analytical results for Zn in OREAS 284 (Certified Value 82 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 85.5185 | 82.4202 | 88 | 83 | 85.0109 | 79 | 73 | 87.2 | 81 | 89 | 77 | 77 | 82 | 83.0947 | 79 | 81 | 79 | 83.3 | 86 | 75 | 83 | 79 | 97 | 122 | 84 |  | 1
| 1 | 2 | 83.723 | 82.9134 | 87 | 83 | 79.0332 | 79 | 75 | 86.8 | 82 | 88 | 76 | 81 | 86 | 80.2961 | 81 | 82 | 80 | 83.3 | 79 | 79 | 83 | 82 | 98 | 111 | 83 |  | 17
| 1 | 3 | 85.4345 | 84.0027 | 88 | 85 | 77.6644 | 79 | 78 | 87.2 | 80 | 87 | 76 | 80 | 84 | 80.5949 | 79 | 80 | 82 | 83 | 82 | 76 | 81 | 82 | 95 | 106 | 82 |  | 16
| 1 | 4 | 85.9115 | 84.5795 | 88 | 88 | 79.4124 | 79 | 81 | 82 | 81 | 88 | 78 | 82 | 84 | 84.9223 | 78 | 81 | 80 | 85.3 | 80 | 79 | 81 | 82 | 94 | 110 | 85 |  | 82.0751
| 1 | 5 | 85.3245 | 82.2568 | 86 | 84 | 74.8425 | 79 | 79 | 85.1 | 83 | 89 | 78 | 81 | 84 | 82.7343 | 80 | 82 | 81 | 86.8 | 82 | 78 | 84 | 81 | 96 | 106 | 83 |  | 70
| 1 | 6 | 86.0645 | 82.6624 | 86 | 86 | 81.7595 | 78 | 81 | 82.6 | 82 | 89 | 75 | 80 | 85 | 83.2006 | 80 | 83 | 84 | 83.5 | 79 | 79 | 82 | 83 | 98 | 109 | 84
| Mean |  | 85.3294 | 83.1392 | 87.1667 | 84.8333 | 79.6205 | 78.8333 | 77.8333 | 85.15 | 81.5 | 88.3333 | 76.6667 | 80.1667 | 84.1667 | 82.4738 | 79.5 | 81.5 | 81 | 84.2 | 81.3333 | 77.6667 | 82.3333 | 81.5 | 96.3333 | 110.6667 | 83.5
| Median |  | 85.4765 | 82.7879 | 87.5 | 84.5 | 79.2228 | 79 | 78.5 | 85.95 | 81.5 | 88.5 | 76.5 | 80.5 | 84 | 82.9145 | 79.5 | 81.5 | 80.5 | 83.4 | 81 | 78.5 | 82.5 | 82 | 96.5 | 109.5 | 83.5
| Std Dev. |  | 0.83744 | 0.93742 | 0.98319 | 1.9408 | 3.4835 | 0.40825 | 3.2506 | 2.3476 | 1.0488 | 0.8165 | 1.2111 | 1.7224 | 1.3292 | 1.7464 | 1.0488 | 1.0488 | 1.7889 | 1.5179 | 2.6583 | 1.7512 | 1.2111 | 1.3784 | 1.633 | 5.9217 | 1.0488
| Rel.Std.Dev. |  | 0.9814% | 1.128% | 1.128% | 2.288% | 4.375% | 0.5179% | 4.176% | 2.757% | 1.287% | 0.9243% | 1.58% | 2.149% | 1.579% | 2.118% | 1.319% | 1.287% | 2.208% | 1.803% | 3.268% | 2.255% | 1.471% | 1.691% | 1.695% | 5.351% | 1.256%
| PDM3 |  | 3.965% | 1.297% | 6.204% | 3.361% | -2.991% | -3.95% | -5.168% | 3.747% | -0.7006% | 7.625% | -6.59% | -2.325% | 2.548% | 0.4858% | -3.137% | -0.7006% | -1.31% | 2.589% | -0.9037% | -5.371% | 0.3147% | -0.7006% | 17.37% | 34.84% | 1.736%
| Z-Score (Absolute) |  | 0.81 | 0.22 | 1.3 | 0.67 | 0.73 | 0.94 | 1.21 | 0.76 | 0.22 | 1.62 | 1.53 | 0.58 | 0.49 | 0.04 | 0.76 | 0.22 | 0.36 | 0.5 | 0.27 | 1.26 | 0 | 0.22 | 3.78 | 7.64 | 0.31
| Analytical results for Zr in OREAS 284 (Certified Value 59 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 62.75 | 57.4852 | 62.4 | 80 | 56.7083 | 51 | 47.6 | 48.7 | 63.8 | 63 | 64.9 | 57.9 | 65.1 | 64.5167 | 59 | 73.5 | 49 | 55.9 | 60 | 55 | 65.8 | 61.5 | 56 | 57.5 | 59.7 |  | 1
| 1 | 2 | 66.81 | 59.8633 | 71.1 | 73 | 53.4304 | 51 | 49 | 58 | 64.5 | 62.1 | 58.9 | 62.4 | 66.6 | 64.815 | 59.6 | 77.6 | 49.8 | 56.3 | 53.2 | 59.2 | 58.7 | 60.8 | 56 | 58.6 | 59 |  | 18
| 1 | 3 | 68.1 | 59.2316 | 61.4 | 78 | 56.8828 | 49 | 49 | 56.5 | 63.2 | 59.8 | 60.9 | 61.4 | 64.1 | 56.0477 | 59.8 | 76 | 49.6 | 60.6 | 56.3 | 56.3 | 58.9 | 63.1 | 54 | 60.2 | 60.4 |  | 16
| 1 | 4 | 65.84 | 59.1992 | 72.8 | 62 | 58.205 | 49 | 49.9 | 49.8 | 64.9 | 62.2 | 65.4 | 61 | 68.4 | 62.4419 | 61.3 | 74.1 | 49.7 | 59.5 | 53.8 | 59.1 | 59 | 57.5 | 54 | 59.4 | 58.6 |  | 58.5493
| 1 | 5 | 64.99 | 56.6327 | 57.3 | 76 | 54.1807 | 50 | 49.8 | 47.8 | 67.5 | 64.2 | 62.4 | 57.2 | 66.4 | 56.7999 | 56.9 | 75.8 | 48.9 | 57.5 | 54.6 | 57.8 | 65.5 | 60.1 | 56 | 59.3 | 57.9 |  | 71
| 1 | 6 | 63.65 | 58.2779 | 59.5 | 64 | 56.0547 | 49 | 51.1 | 46.9 | 67.2 | 62.5 | 62 | 60.4 | 63.1 | 59.436 | 59.1 | 76.1 | 46 | 62.1 | 53.4 | 58 | 64.5 | 65 | 56 | 60.2 | 58.3
| Mean |  | 65.3567 | 58.4483 | 64.0833 | 72.1667 | 55.9103 | 49.8333 | 49.4 | 51.2833 | 65.1833 | 62.3 | 62.4167 | 60.05 | 65.6167 | 60.6762 | 59.2833 | 75.5167 | 48.8333 | 58.65 | 55.2167 | 57.5667 | 62.0667 | 61.3333 | 55.3333 | 59.2 | 58.9833
| Median |  | 65.415 | 58.7386 | 61.9 | 74.5 | 56.3815 | 49.5 | 49.4 | 49.25 | 64.7 | 62.35 | 62.2 | 60.7 | 65.75 | 60.9389 | 59.35 | 75.9 | 49.3 | 58.5 | 54.2 | 57.9 | 61.75 | 61.15 | 56 | 59.35 | 58.8
| Std Dev. |  | 1.9856 | 1.2193 | 6.3603 | 7.4944 | 1.7894 | 0.98319 | 1.1713 | 4.7444 | 1.7792 | 1.445 | 2.4457 | 2.055 | 1.9094 | 3.8206 | 1.4303 | 1.4878 | 1.4376 | 2.4833 | 2.6003 | 1.6403 | 3.5331 | 2.5727 | 1.0328 | 1.0296 | 0.92826
| Rel.Std.Dev. |  | 3.038% | 2.086% | 9.925% | 10.38% | 3.201% | 1.973% | 2.371% | 9.251% | 2.73% | 2.319% | 3.918% | 3.422% | 2.91% | 6.297% | 2.413% | 1.97% | 2.944% | 4.234% | 4.709% | 2.849% | 5.692% | 4.195% | 1.866% | 1.739% | 1.574%
| PDM3 |  | 11.63% | -0.1725% | 9.452% | 23.26% | -4.507% | -14.89% | -15.63% | -12.41% | 11.33% | 6.406% | 6.605% | 2.563% | 12.07% | 3.633% | 1.254% | 28.98% | -16.59% | 0.1719% | -5.692% | -1.678% | 6.007% | 4.755% | -5.493% | 1.111% | 0.7412%
| Z-Score (Absolute) |  | 1.21 | 0.17 | 0.96 | 2.58 | 0.67 | 1.89 | 1.98 | 1.6 | 1.18 | 0.6 | 0.63 | 0.15 | 1.27 | 0.28 | 0 | 3.25 | 2.09 | 0.13 | 0.81 | 0.34 | 0.56 | 0.41 | 0.79 | 0.02 | 0.06
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C973: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C991: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1010: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1047: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1065: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1084: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 284 (Certified Value 0.082 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | < 0.1 | 0.08 | 0.49774 | < 1 | 0.092 |  | < 0.5 | < 0.3 | 0.052 | 0.09 | 0.08 | 0.13 | 0.08 | 0.08 | 0.088027 | 0.1 | 0.07 | < 0.05 | 0.09 | 0.078 | 0.08 | 0.06 | 0.09 | < 2 | 0.241 | 0.07 |  |  |  | 1
| 1 | 2 | < 0.1 | 0.09 | 0.43813 | < 1 | 0.084 | 3.3535 | < 0.5 | < 0.3 | 0.062 | 0.1 | 0.09 | 0.08 | 0.08 | 0.09 | 0.091274 | < 0.1 | 0.08 | < 0.05 | 0.08 | 0.078 | 0.08 | 0.06 | 0.09 | < 2 | 0.24 | 0.09 |  |  |  | 22
| 1 | 3 | < 0.1 | 0.1 | 0.71841 | < 1 | 0.094 |  | < 0.5 | < 0.3 | 0.102 | 0.08 | 0.07 | 0.08 | 0.07 | 0.08 | 0.090841 | < 0.1 | 0.08 | < 0.05 | 0.08 | 0.076 | 0.09 | 0.06 | 0.09 | < 2 | 0.23 | 0.08 |  |  |  | 16
| 1 | 4 | < 0.1 | 0.09 | 0.332 | < 1 | 0.083 |  | < 0.5 | < 0.3 | 0.062 | 0.09 | 0.08 | 0.06 | 0.08 | 0.07 | 0.084645 | < 0.1 | 0.08 | 0.05 | 0.08 | 0.08 | 0.08 | 0.06 | 0.09 | < 2 | 0.23 | 0.08 |  |  |  | 0.082347
| 1 | 5 | < 0.1 | 0.08 | 0.48467 | < 1 | 0.093 |  | < 0.5 | < 0.3 | 0.052 | 0.09 | 0.09 | 0.08 | 0.08 | 0.08 | 0.085018 | < 0.1 | 0.07 | < 0.05 | 0.08 | 0.081 | 0.08 | 0.06 | 0.08 | < 2 | 0.23 | 0.08 |  |  |  | 73
| 1 | 6 | < 0.1 | 0.08 | 0.34613 | < 1 | 0.084 |  | < 0.5 | < 0.3 | 0.062 | 0.08 | 0.07 | 0.1 | 0.1 | 0.08 | 0.086542 | < 0.1 | 0.08 | 0.05 | 0.08 | 0.076 | 0.08 |  | 0.09 | < 2 | 0.281 | 0.07
| Mean |  |  | 0.086667 | 0.46951 |  | 0.088333 | 3.3535 |  |  | 0.065333 | 0.088333 | 0.08 | 0.088333 | 0.081667 | 0.08 | 0.087725 | 0.1 | 0.076667 | 0.05 | 0.081667 | 0.078167 | 0.081667 | 0.06 | 0.088333 |  | 0.242 | 0.078333
| Median |  |  | 0.085 | 0.4614 |  | 0.088 | 3.3535 |  |  | 0.062 | 0.09 | 0.08 | 0.08 | 0.08 | 0.08 | 0.087285 | 0.1 | 0.08 | 0.05 | 0.08 | 0.078 | 0.08 | 0.06 | 0.09 |  | 0.235 | 0.08
| Std Dev. |  |  | 0.008165 | 0.14007 |  | 0.005164 |  |  |  | 0.018619 | 0.0075277 | 0.0089443 | 0.024014 | 0.0098319 | 0.0063246 | 0.0028497 |  | 0.005164 | 0 | 0.0040825 | 0.0020412 | 0.0040825 | 0 | 0.0040825 |  | 0.019789 | 0.0075277
| Rel.Std.Dev. |  |  | 9.421% | 29.83% |  | 5.846% |  |  |  | 28.5% | 8.522% | 11.18% | 27.19% | 12.04% | 7.906% | 3.248% |  | 6.736% | 0% | 4.999% | 2.611% | 4.999% | 0% | 4.622% |  | 8.177% | 9.61%
| PDM3 |  |  | 5.245% | 470.2% |  | 7.269% | 3972% |  |  | -20.66% | 7.269% | -2.851% | 7.269% | -0.8266% | -2.851% | 6.53% | 21.44% | -6.898% | -39.28% | -0.8266% | -5.077% | -0.8266% | -27.14% | 7.269% |  | 193.9% | -4.875%
| Z-Score (Absolute) |  | 3.41 | 0.25 | 38.5 | 41.54 | 0.42 | 326.6 | 16.57 | 6.58 | 1.88 | 0.42 | 0.42 | 0.42 | 0.25 | 0.42 | 0.36 | 2.58 | 0.75 | 5.08 | 0.25 | 0.6 | 0.25 | 2.41 | 0.42 | 91.49 | 15.77 | 0.58
| Analytical results for Al in OREAS 284 (Certified Value 2.16 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 2.1093 | 2.1217 | 2.2883 | 2.23 | 2.376 | 2.25 | 1.91 | 2.2 | 2.47 | 2.07 | 2.17 | 2.14 | 2.087 | 2.01 | 2.08 | 2.32 | 1.8 | 2.08 | 2.06 | 2.0304 | 1.8944 | 2.23 | 2.228 | 2.489 | 2.54 | 2.21 |  |  |  | 1
| 1 | 2 | 2.0659 | 2.1064 | 2.2471 | 2.14 | 2.5464 | 2.27 | 1.99 | 2.22 | 2.46 | 2.09 | 2.12 | 2.11 | 2.1092 | 2.01 | 2.09 | 2.29 | 1.82 | 2.1 | 1.98 | 1.981 | 2.0227 | 2.186 | 2.138 | 2.623 | 2.39 | 2.26 |  |  |  | 0
| 1 | 3 | 2.0439 | 2.078 | 2.325 | 2.2 | 2.2811 | 2.26 | 1.98 | 2.22 | 2.43 | 2.04 | 2.16 | 2.11 | 2.1007 | 2.03 | 2.04 | 2.29 | 1.8 | 1.89 | 2.02 | 2.0385 | 2.0749 | 2.405 | 2.143 | 2.389 | 2.31 | 2.28 |  |  |  | 16
| 1 | 4 | 2.0363 | 2.1223 | 2.2582 | 2.18 | 2.2345 | 2.23 | 2.02 | 2.21 | 2.42 | 2.02 | 2.19 | 2.12 | 2.1751 | 2.05 | 2.09 | 2.35 | 1.77 | 2.08 | 2.06 | 1.9757 | 2.0231 | 2.202 | 2.186 | 2.252 | 2.46 | 2.28 |  |  |  | 2.1645
| 1 | 5 | 2.0738 | 2.0609 | 2.3107 | 2.15 | 2.3123 | 2.24 | 2 | 2.25 | 2.42 | 2.07 | 2.19 | 2.16 | 2.1351 | 2.07 | 2.07 | 2.29 | 1.82 | 2.11 | 2.05 | 2.0109 | 2.0903 | 2.541 | 2.17 | 2.392 | 2.09 | 2.29 |  |  |  | 74
| 1 | 6 | 2.1216 | 2.1568 | 2.3434 | 1.86 | 2.3474 | 2.24 | 1.88 | 2.01 | 2.21 | 2.07 | 2.17 | 2.23 | 2.1705 | 2.09 | 2.05 | 2.26 | 1.79 | 2.1 | 2.05 | 1.9304 | 2.0124 | 2.392 | 2.149 | 2.461 | 2.24 | 2.33
| Mean |  | 2.0751 | 2.1077 | 2.2954 | 2.1267 | 2.3496 | 2.2483 | 1.9633 | 2.185 | 2.4017 | 2.06 | 2.1667 | 2.145 | 2.1296 | 2.0433 | 2.07 | 2.3 | 1.8 | 2.06 | 2.0367 | 1.9945 | 2.0196 | 2.326 | 2.169 | 2.4343 | 2.3383 | 2.275
| Median |  | 2.0699 | 2.1141 | 2.2995 | 2.165 | 2.3298 | 2.245 | 1.985 | 2.215 | 2.425 | 2.07 | 2.17 | 2.13 | 2.1222 | 2.04 | 2.075 | 2.29 | 1.8 | 2.09 | 2.05 | 1.9959 | 2.0229 | 2.311 | 2.1595 | 2.4265 | 2.35 | 2.28
| Std Dev. |  | 0.034362 | 0.034334 | 0.037902 | 0.13471 | 0.10841 | 0.01472 | 0.055377 | 0.08735 | 0.096212 | 0.025298 | 0.02582 | 0.045935 | 0.036984 | 0.03266 | 0.020976 | 0.030984 | 0.018974 | 0.084143 | 0.031411 | 0.040344 | 0.069 | 0.14214 | 0.034082 | 0.1236 | 0.16142 | 0.03937
| Rel.Std.Dev. |  | 1.656% | 1.629% | 1.651% | 6.334% | 4.614% | 0.6547% | 2.821% | 3.998% | 4.006% | 1.228% | 1.192% | 2.141% | 1.737% | 1.598% | 1.013% | 1.347% | 1.054% | 4.085% | 1.542% | 2.023% | 3.416% | 6.111% | 1.571% | 5.077% | 6.903% | 1.731%
| PDM3 |  | -4.127% | -2.623% | 6.052% | -1.746% | 8.554% | 3.875% | -9.292% | 0.949% | 10.96% | -4.826% | 0.1019% | -0.8991% | -1.609% | -5.596% | -4.364% | 6.262% | -16.84% | -4.826% | -5.904% | -7.853% | -6.691% | 7.463% | 0.2097% | 12.47% | 8.033% | 5.107%
| Z-Score (Absolute) |  | 0.4 | 0.19 | 1.01 | 0.07 | 1.35 | 0.71 | 1.11 | 0.3 | 1.68 | 0.49 | 0.19 | 0.05 | 0.05 | 0.6 | 0.43 | 1.04 | 2.15 | 0.49 | 0.64 | 0.91 | 0.75 | 1.2 | 0.2 | 1.89 | 1.28 | 0.88
| Analytical results for As in OREAS 284 (Certified Value 174 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | 0
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0
| 1 | 1 | 166.96 | 177.724 | 184 | 189.7746 | 158.4207 | 200 | 183.306 | 172 | 172.8 | 169.7 | 182 | 194 | 174.5 | 179 | 171.5 | 170.057 | 170.9 | 159 | 173 | 177 | 174.4 | 183 | 169 | 164 | 173.28 | 159 | 175.4 | 195 | 166 | 1
| 1 | 2 | 163.69 | 173.7796 | 185 | 183.677 | 152.006 | 190 | 196.6906 | 173 | 166.7 | 179.6 | 184.5 | 193 | 175 | 177.5 | 174.5 | 172.8309 | 170.6 | 164 | 172 | 175 | 182.3 | 173 | 168 | 173 | 172.88 | 159 | 183.1 | 194 | 168 | 12
| 1 | 3 | 162.21 | 175.7684 | 186 | 185.0068 | 153.2494 | 192 | 166.4967 | 173 | 169.2 | 178.4 | 185 | 183 | 172 | 180.5 | 171.5 | 169.9476 | 171.1 | 163 | 174 | 174 | 177.3 | 175 | 172 | 179 | 173.64 | 162 | 166.3 | 190 | 174 | 16
| 1 | 4 | 165.42 | 174.8903 | 183 | 173.2203 | 151.3022 | 194 | 155.0633 | 172 | 165.5 | 179.3 | 184.5 | 191.5 | 171.5 | 183 | 169 | 171.505 | 168.8 | 166 | 172 | 176 | 184.5 | 175 | 166 | 174 | 168.46 | 165 | 159.5 | 178 | 173 | 174.1327
| 1 | 5 | 165.55 | 176.8692 | 183 | 184.6729 | 155.5944 | 200 | 166.3245 | 173 | 167.4 | 177.3 | 186 | 192.5 | 177.5 | 182 | 175 | 169.4975 | 169.4 | 161 | 172 | 176 | 183.7 | 181 | 167 | 181 | 177.2 | 161 | 165.4 | 185 | 170 | 75
| 1 | 6 | 163.82 | 177.5771 | 184 | 194.8623 | 153.8869 | 186 | 167.4159 | 174 | 168.7 | 169.9 | 178 | 187 | 174 | 179.5 | 177.5 | 172.4049 | 172 | 165 | 174 | 174 | 178.8 | 179 | 165 | 173 | 168.53 | 163 | 170.3 | 190 | 176
| Mean |  | 164.6083 | 176.1014 | 184.1667 | 185.2023 | 154.0766 | 193.6667 | 172.5495 | 172.8333 | 168.3833 | 175.7 | 183.3333 | 190.1667 | 174.0833 | 180.25 | 173.1667 | 171.0405 | 170.4667 | 163 | 172.8333 | 175.3333 | 180.1667 | 177.6667 | 167.8333 | 174 | 172.3317 | 161.5 | 170 | 188.6667 | 171.1667
| Median |  | 164.62 | 176.3188 | 184 | 184.8399 | 153.5682 | 193 | 166.9563 | 173 | 168.05 | 177.85 | 184.5 | 192 | 174.25 | 180 | 173 | 170.781 | 170.75 | 163.5 | 172.5 | 175.5 | 180.55 | 177 | 167.5 | 173.5 | 173.08 | 161.5 | 168.3 | 190 | 171.5
| Std Dev. |  | 1.6907 | 1.5728 | 1.169 | 7.2171 | 2.6031 | 5.5737 | 14.8728 | 0.75277 | 2.5451 | 4.6403 | 2.9269 | 4.274 | 2.1775 | 2.0187 | 3.0605 | 1.4018 | 1.1725 | 2.6077 | 0.98319 | 1.2111 | 3.9788 | 3.9328 | 2.4833 | 5.933 | 3.3483 | 2.3452 | 8.3182 | 6.314 | 3.8166
| Rel.Std.Dev. |  | 1.027% | 0.8931% | 0.6348% | 3.897% | 1.689% | 2.878% | 8.619% | 0.4355% | 1.512% | 2.641% | 1.596% | 2.247% | 1.251% | 1.12% | 1.767% | 0.8196% | 0.6878% | 1.6% | 0.5689% | 0.6907% | 2.208% | 2.214% | 1.48% | 3.41% | 1.943% | 1.452% | 4.893% | 3.347% | 2.23%
| PDM3 |  | -5.47% | 1.131% | 5.762% | 6.357% | -11.52% | 11.22% | -0.9092% | -0.7462% | -3.302% | 0.9001% | 5.284% | 9.208% | -0.02835% | 3.513% | -0.5548% | -1.776% | -2.105% | -6.393% | -0.7462% | 0.6895% | 3.465% | 2.029% | -3.618% | -0.07621% | -1.034% | -7.255% | -2.373% | 8.346% | -1.703%
| Z-Score (Absolute) |  | 1.28 | 0.44 | 1.65 | 1.8 | 2.86 | 3.07 | 0.09 | 0.05 | 0.72 | 0.38 | 1.52 | 2.55 | 0.14 | 1.06 | 0 | 0.32 | 0.4 | 1.52 | 0.05 | 0.32 | 1.05 | 0.67 | 0.8 | 0.12 | 0.13 | 1.75 | 0.47 | 2.32 | 0.3
| Analytical results for B in OREAS 284 (Indicative Value 9.78 ppm) | [30 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.26 | 1.28 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | [11 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 15g | 01g | 0.5g | 0.25g | [11 empty] | 2
| 1 | 1 | 4.2953 | < 10 | 3 | 3.6243 | 15 | < 20 | < 10 | < 10 | < 10 | 10 | < 10 | < 10 | < 10 | 17 | 2 | < 10 | 26.16 | < 10 | [11 empty] | 1
| 1 | 2 | 4.3592 | < 10 | 3 | 6.4351 | 14 | < 20 | < 10 | < 10 | < 10 | 10 | < 10 | < 10 | < 10 | 18 | 3 | < 10 | 24.28 | < 10 | [11 empty] | 1
| 1 | 3 | 4.2756 | < 10 | 2 | 2.4915 | 13 | < 20 | < 10 | 10 | < 10 | 10 | < 10 | < 10 | < 10 | 19 | 2 | < 10 | 22.08 | < 10 | [11 empty] | 16
| 1 | 4 | 4.1978 | < 10 | 2 | 4.0473 | 12 | < 20 | < 10 | 10 | < 10 | 10 | < 10 | < 10 | < 10 | 17 | 2 | < 10 | 27.48 | < 10 | [11 empty] | 9.7769
| 1 | 5 | 1.02: 4.1582; 1.03: < 10; 1.04: 3; 1.05: 1.653; 1.07: 13; 1.09: < 20; 1.10: < 10; 1.11: 10; 1.12: < 10; 1.13: 10; 1.14: < 10; 1.15: < 10; 1.16: < 10; 1.17: 18; 1.18: 2; 1.19: < 10; 1.21: < 10 | 10
| 1 | 6 | 4.0703 | < 10 | 2 | 3.6994 | 14 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 17 | 2 | < 10 | 21.37 | < 10
| Mean |  | 1.02: 4.2261; 1.04: 2.5; 1.05: 3.6584; 1.07: 13.5; 1.11: 10; 1.13: 10; 1.17: 17.6667; 1.18: 2.1667; 1.20: 24.274
| Median |  | 1.02: 4.2367; 1.04: 2.5; 1.05: 3.6618; 1.07: 13.5; 1.11: 10; 1.13: 10; 1.17: 17.5; 1.18: 2; 1.20: 24.28
| Std Dev. |  | 1.02: 0.10455; 1.04: 0.54772; 1.05: 1.6277; 1.07: 1.0488; 1.11: 0; 1.13: 0; 1.17: 0.8165; 1.18: 0.40825; 1.20: 2.602
| Rel.Std.Dev. |  | 1.02: 2.474%; 1.04: 21.91%; 1.05: 44.49%; 1.07: 7.769%; 1.11: 0%; 1.13: 0%; 1.17: 4.622%; 1.18: 18.84%; 1.20: 10.72%
| PDM3 |  | 1.02: -56.77%; 1.04: -74.43%; 1.05: -62.58%; 1.07: 38.08%; 1.11: 2.282%; 1.13: 2.282%; 1.17: 80.7%; 1.18: -77.84%; 1.20: 148.3%
| Z-Score (Absolute) |  | 0.49 | 0 | 1.59 | 0.86 | 5.42 | 3.19 | 0 | 1.59 | 0 | 2.66 | 0 | 0 | 0 | 8.07 | 1.81 | 0 | 12.29 | 0
| Analytical results for Ba in OREAS 284 (Certified Value 510 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 0
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 0
| 1 | 1 | 501.2011 | 529 | 607.2657 | 521.3701 | 336 | 526.3062 | 486 | 497.3 | 540 | 570 | 510 | 520 | 500 | 467.8977 | 517 | 452 | 530 | 497 | 423.9 | 505 | 484 | 519 | 455.41 | 358 | 548.1 | 553 | 483 |  |  | 1
| 1 | 2 | 499.6098 | 528 | 599.8032 | 490.8273 | 268 | 558.5518 | 489 | 515.6 | 540 | 580 | 510 | 510 | 500 | 491.8727 | 545 | 483 | 531 | 493 | 453.9 | 486 | 475 | 543 | 470.51 | 351 | 545.49 | 530 | 494 |  |  | 13
| 1 | 3 | 504.1546 | 525 | 577.7881 | 498.2358 | 246 | 496.3855 | 487 | 511 | 540 | 570 | 500 | 520 | 500 | 480.9297 | 545 | 473 | 533 | 504 | 423.3 | 497 | 486 | 554 | 474.32 | 372 | 536.31 | 510 | 496 |  |  | 16
| 1 | 4 | 510.9042 | 511 | 562.1547 | 514.0349 | 301 | 490.6827 | 477 | 518.6 | 540 | 560 | 500 | 520 | 500 | 479.0621 | 554 | 469 | 527 | 504 | 429.8 | 504 | 468 | 539 | 471.94 | 344 | 506.61 | 540 | 499 |  |  | 510.0781
| 1 | 5 | 504.3843 | 518 | 583.734 | 505.7419 | 334 | 499.5736 | 485 | 521.4 | 540 | 570 | 510 | 520 | 510 | 490.6313 | 557 | 461 | 527 | 502 | 436.5 | 499 | 478 | 565 |  | 324 | 539.46 | 480 | 501 |  |  | 76
| 1 | 6 | 511.7546 | 530 | 631.6734 | 501.6472 | 297 | 508.1042 | 480 | 499.1 | 510 | 510 | 510 | 510 | 530 | 469.5209 | 563 | 468 | 530 | 508 | 454.4 | 501 | 460 | 536 | 455.68 | 340 | 528.57 | 519 | 509
| Mean |  | 505.3348 | 523.5 | 593.7365 | 505.3095 | 297 | 513.2673 | 484 | 510.5 | 535 | 560 | 506.6667 | 516.6667 | 506.6667 | 479.9857 | 546.8333 | 467.6667 | 529.6667 | 501.3333 | 436.9667 | 498.6667 | 475.1667 | 542.6667 | 465.572 | 348.1667 | 534.09 | 522 | 497
| Median |  | 504.2695 | 526.5 | 591.7686 | 503.6945 | 299 | 503.8389 | 485.5 | 513.3 | 540 | 570 | 510 | 520 | 500 | 479.9959 | 549.5 | 468.5 | 530 | 503 | 433.15 | 500 | 476.5 | 541 | 470.51 | 347.5 | 537.885 | 524.5 | 497.5
| Std Dev. |  | 4.9878 | 7.5033 | 24.5392 | 11.0222 | 35.6314 | 25.4163 | 4.5607 | 10.1477 | 12.2474 | 25.2982 | 5.164 | 5.164 | 12.1106 | 10.1194 | 16.2039 | 10.5388 | 2.3381 | 5.4283 | 14.1404 | 6.8896 | 9.8472 | 15.7818 | 9.2545 | 16.3758 | 15.1369 | 25.5578 | 8.6023
| Rel.Std.Dev. |  | 0.987% | 1.433% | 4.133% | 2.181% | 12% | 4.952% | 0.9423% | 1.988% | 2.289% | 4.518% | 1.019% | 0.9995% | 2.39% | 2.108% | 2.963% | 2.253% | 0.4414% | 1.083% | 3.236% | 1.382% | 2.072% | 2.908% | 1.988% | 4.703% | 2.834% | 4.896% | 1.731%
| PDM3 |  | -0.9299% | 2.631% | 16.4% | -0.9349% | -41.77% | 0.6253% | -5.113% | 0.08272% | 4.886% | 9.787% | -0.6688% | 1.292% | -0.6688% | -5.9% | 7.206% | -8.315% | 3.84% | -1.714% | -14.33% | -2.237% | -6.844% | 6.389% | -8.725% | -31.74% | 4.708% | 2.337% | -2.564%
| Z-Score (Absolute) |  | 0.04 | 0.49 | 2.55 | 0.04 | 6.15 | 0.19 | 0.66 | 0.11 | 0.83 | 1.56 | 0 | 0.29 | 0 | 0.78 | 1.18 | 1.14 | 0.67 | 0.16 | 2.04 | 0.23 | 0.92 | 1.06 | 1.2 | 4.65 | 0.8 | 0.45 | 0.28
| Analytical results for Be in OREAS 284 (Certified Value 1.66 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 1.7288 | 1.7 | 2.062 | 2.0836 | 1.6 | 1.7056 | < 5 | 1.67 | 1.75 | 1.55 | 1.66 | 1.64 | 1.4632 | 1 | 1.6 | 1.8 | 1.51 | 1.8 | 1.62 | 1.26 | 1.58 | 1.7 | [7 empty] | 1
| 1 | 2 | 1.7157 | 1.69 | 2.1159 | 1.985 | 1.6 | 1.7868 | < 5 | 1.68 | 1.76 | 1.57 | 1.64 | 1.63 | 1.5131 | 2 | 1.7 | 1.8 | 1.68 | 1.7 | 1.58 | 1.22 | 1.55 | 1.7 | [7 empty] | 25
| 1 | 3 | 1.6034 | 1.7 | 2.054 | 2.1654 | 1.7 | 1.7026 | < 5 | 1.69 | 1.7 | 1.53 | 1.68 | 1.61 | 1.4789 | 1 | 1.7 | 1.8 | 1.64 | 1.8 | 1.69 | 1.22 | 1.2 | 1.7 | [7 empty] | 16
| 1 | 4 | 1.5805 | 1.68 | 1.9588 | 1.9445 | 1.7 | 1.7517 | < 5 | 1.68 | 1.71 | 1.5 | 1.7 | 1.63 | 1.5869 | 2 | 1.7 | 1.8 | 1.48 | 1.7 | 1.69 | 1.25 | 1.56 | 1.6 | [7 empty] | 1.6593
| 1 | 5 | 1.7177 | 1.69 | 2.1057 | 2.0613 | 1.8 | 1.7031 | < 5 | 1.71 | 1.74 | 1.56 | 1.7 | 1.66 | 1.4889 | 1 | 1.6 | 1.8 | 1.4 | 1.8 | 1.85 | 1.23 | 1.7 | 1.7 | [7 empty] | 77
| 1 | 6 | 1.5792 | 1.66 | 2.1222 | 2.0752 | 1.7 | 1.6727 | < 5 | 1.52 | 1.59 | 1.55 | 1.66 | 1.7 | 1.6113 | 2 | 1.7 | 1.7 | 1.47 | 1.8 | 1.61 | 1.18 | 1.39 | 1.7
| Mean |  | 1.6542 | 1.6867 | 2.0698 | 2.0525 | 1.6833 | 1.7204 |  | 1.6583 | 1.7083 | 1.5433 | 1.6733 | 1.645 | 1.5237 | 1.5 | 1.6667 | 1.7833 | 1.53 | 1.7667 | 1.6733 | 1.2267 | 1.4967 | 1.6833
| Median |  | 1.6595 | 1.69 | 2.0838 | 2.0683 | 1.7 | 1.7044 |  | 1.68 | 1.725 | 1.55 | 1.67 | 1.635 | 1.501 | 1.5 | 1.7 | 1.8 | 1.495 | 1.8 | 1.655 | 1.225 | 1.555 | 1.7
| Std Dev. |  | 0.073493 | 0.015055 | 0.061293 | 0.078143 | 0.075277 | 0.041229 |  | 0.069113 | 0.062423 | 0.025033 | 0.024221 | 0.031464 | 0.061072 | 0.54772 | 0.05164 | 0.040825 | 0.1077 | 0.05164 | 0.097297 | 0.028048 | 0.1758 | 0.040825
| Rel.Std.Dev. |  | 4.443% | 0.8926% | 2.961% | 3.807% | 4.472% | 2.396% |  | 4.168% | 3.654% | 1.622% | 1.447% | 1.913% | 4.008% | 36.51% | 3.098% | 2.289% | 7.039% | 2.923% | 5.815% | 2.286% | 11.75% | 2.425%
| PDM3 |  | -0.3035% | 1.652% | 24.74% | 23.7% | 1.451% | 3.686% |  | -0.05567% | 2.958% | -6.986% | 0.8483% | -0.8592% | -8.168% | -9.598% | 0.4466% | 7.478% | -7.79% | 6.473% | 0.8483% | -26.07% | -9.799% | 1.451%
| Z-Score (Absolute) |  | 0.27 | 0.19 | 5.68 | 5.43 | 0.14 | 0.67 | 11.84 | 0.21 | 0.5 | 1.86 | 0 | 0.41 | 2.14 | NA | 0.1 | 1.58 | 2.05 | 1.34 | 0 | 6.4 | 2.53 | 0.14
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 284 (Certified Value 0.69 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.70517 | 0.85 | 0.58756 | 0.71 | < 2 | 0.66 | 0.66 | 0.74 | 0.65 | 0.84 | 0.75 | 0.71 | 0.65116 | 0.6 | 0.63 | 0.73 | 0.74 | 0.68 | 0.85 | < 2 | 0.59 | 0.66 | < 0.5 | 0.706 | 1 |  |  |  |  | 1
| 1 | 2 | 0.63541 | 0.7 | 0.61399 | 0.68 | < 2 | 0.6 | 0.82 | 0.71 | 0.75 | 0.53 | 0.76 | 0.81 | 0.66337 | 0.5 | 0.68 | 0.68 | 0.8 | 0.66 | 0.67 | < 2 | 0.71 | 0.59 | < 0.5 | 0.67 | 0.83 |  |  |  |  | 26
| 1 | 3 | 0.64402 | 0.86 | 0.63738 | 0.99 | < 2 | 0.61 | 1.04 | 0.57 | 0.59 | 0.74 | 0.58 | 0.73 | 0.76894 | 0.5 | 0.67 | 0.68 | 1.89 | 0.64 | 0.5 | < 2 | 0.81 | 0.68 | < 0.5 | 0.7 | 0.63 |  |  |  |  | 16
| 1 | 4 | 0.70634 | 0.77 | 0.79035 | 0.7 | < 2 | 0.93 | 0.65 | 0.64 | 0.64 | 0.7 | 0.57 | 0.71 | 0.72354 | 0.4 | 0.65 | 0.69 | 0.71 | 0.76 | 0.64 | < 2 | 0.68 | 0.55 | < 0.5 | 0.72 | 0.62 |  |  |  |  | 0.69424
| 1 | 5 | 0.65617 | 0.7 | 0.88699 | 0.62 | < 2 | 0.83 | 0.67 | 0.91 | 0.93 | 0.59 | 0.7 | 0.66 | 0.707 | 0.5 | 0.61 | 0.75 | 0.71 | 0.73 | 0.61 | < 2 | 0.76 | 0.5 | < 0.5 | 0.76 | 0.74 |  |  |  |  | 78
| 1 | 6 | 0.67873 | 0.68 | 0.87585 | 0.84 | < 2 | 0.5 | 0.96 | 0.63 | 0.64 | 0.66 | 0.59 | 0.56 | 0.59794 | 0.5 | 0.69 | 0.75 | 0.67 | 0.65 | 0.75 | < 2 | 0.7 | 0.62 | < 0.5 | 0.672 | 1.09
| Mean |  | 0.67097 | 0.76 | 0.73202 | 0.75667 |  | 0.68833 | 0.8 | 0.7 | 0.7 | 0.67667 | 0.65833 | 0.69667 | 0.68533 | 0.5 | 0.655 | 0.71333 | 0.92 | 0.68667 | 0.67 |  | 0.70833 | 0.6 |  | 0.70467 | 0.81833
| Median |  | 0.66745 | 0.735 | 0.71386 | 0.705 |  | 0.635 | 0.745 | 0.675 | 0.645 | 0.68 | 0.645 | 0.71 | 0.68519 | 0.5 | 0.66 | 0.71 | 0.725 | 0.67 | 0.655 |  | 0.705 | 0.605 |  | 0.703 | 0.785
| Std Dev. |  | 0.030627 | 0.07975 | 0.13554 | 0.13515 |  | 0.16043 | 0.16888 | 0.11933 | 0.12426 | 0.11003 | 0.088412 | 0.082865 | 0.060344 | 0.063246 | 0.030822 | 0.033862 | 0.47716 | 0.048028 | 0.12017 |  | 0.074677 | 0.067823 |  | 0.033435 | 0.19385
| Rel.Std.Dev. |  | 4.565% | 10.49% | 18.52% | 17.86% |  | 23.31% | 21.11% | 17.05% | 17.75% | 16.26% | 13.43% | 11.89% | 8.805% | 12.65% | 4.706% | 4.747% | 51.87% | 6.994% | 17.94% |  | 10.54% | 11.3% |  | 4.745% | 23.69%
| PDM3 |  | -3.351% | 9.472% | 5.442% | 8.992% |  | -0.8508% | 15.23% | 0.8297% | 0.8297% | -2.531% | -5.172% | 0.3495% | -1.284% | -27.98% | -5.652% | 2.75% | 32.52% | -1.091% | -3.492% |  | 2.03% | -13.57% |  | 1.502% | 17.87%
| Z-Score (Absolute) |  | 0.63 | 1.3 | 0.7 | 1.23 | 6.52 | 0.25 | 2.17 | 0 | 0 | 0.51 | 0.91 | 0.07 | 0.32 | NA | 0.98 | 0.29 | 4.78 | 0.29 | 0.65 | 6.52 | 0.18 | 2.17 | 9.79 | 0.1 | 2.57
Results from laboratory 1.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 284 (Certified Value 0.572 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.57706 | 0.55 | 0.59086 | 0.59859 | 0.72 | 0.7521 | 0.59 | 0.432 | 0.57349 | 0.59 | 0.64 | 0.56 | 0.57 | 0.57 | 0.55222 | 0.56 | 0.56 | 0.58 | 0.52 | 0.55 | 0.54 | 0.54 | 0.49 | 0.611 | 0.615 | < 2 | 0.631 | 0.59 |  | 1
| 1 | 2 | 0.5513 | 0.54 | 0.59524 | 0.58925 | 0.71 | 0.8245 | 0.6 | 0.4364 | 0.61461 | 0.59 | 0.64 | 0.57 | 0.56 | 0.57 | 0.55594 | 0.56 | 0.56 | 0.59 | 0.52 | 0.56 | 0.52 | 0.53 | 0.52 | 0.578 | 0.586 | < 2 | 0.62 | 0.6 |  | 0
| 1 | 3 | 0.55118 | 0.54 | 0.57862 | 0.61333 | 0.7 | 0.7227 | 0.59 | 0.4431 | 0.61281 | 0.59 | 0.63 | 0.56 | 0.57 | 0.57 | 0.49601 | 0.56 | 0.55 | 0.59 | 0.52 | 0.51 | 0.53 | 0.54 | 0.54 |  | 0.593 | < 2 | 0.59 | 0.6 |  | 16
| 1 | 4 | 0.55735 | 0.55 | 0.56255 | 0.59678 | 0.68 | 0.7186 | 0.59 | 0.4327 | 0.61301 | 0.58 | 0.63 | 0.55 | 0.58 | 0.57 | 0.56815 | 0.57 | 0.57 | 0.58 | 0.51 | 0.55 | 0.54 | 0.52 | 0.52 | 0.57 | 0.607 | < 2 | 0.57 | 0.61 |  | 0.57184
| 1 | 5 | 0.56176 | 0.53 | 0.6014 | 0.61066 | 0.73 | 0.7302 | 0.59 | 0.4313 | 0.61301 | 0.6 | 0.63 | 0.56 | 0.58 | 0.58 | 0.52263 | 0.58 | 0.56 | 0.57 | 0.52 | 0.56 | 0.54 | 0.54 | 0.54 | 0.655 | 0.6 | < 2 | 0.57 | 0.61 |  | 79
| 1 | 6 | 0.5676 | 0.55 | 0.57481 | 0.61738 | 0.65 | 0.7692 | 0.59 | 0.4341 | 0.57229 | 0.54 | 0.59 | 0.56 | 0.57 | 0.59 | 0.49185 | 0.58 | 0.55 | 0.58 | 0.51 | 0.56 | 0.54 | 0.52 | 0.53 | 0.607 | 0.593 | < 2 | 0.612 | 0.62
| Mean |  | 0.56104 | 0.54333 | 0.58391 | 0.60433 | 0.69833 | 0.75288 | 0.59167 | 0.43493 | 0.59987 | 0.58167 | 0.62667 | 0.56 | 0.57167 | 0.575 | 0.53113 | 0.56833 | 0.55833 | 0.58167 | 0.51667 | 0.54833 | 0.535 | 0.53167 | 0.52333 | 0.6042 | 0.599 |  | 0.59883 | 0.605
| Median |  | 0.55956 | 0.545 | 0.58474 | 0.60462 | 0.705 | 0.74115 | 0.59 | 0.4334 | 0.61291 | 0.59 | 0.63 | 0.56 | 0.57 | 0.57 | 0.53743 | 0.565 | 0.56 | 0.58 | 0.52 | 0.555 | 0.54 | 0.535 | 0.525 | 0.607 | 0.5965 |  | 0.601 | 0.605
| Std Dev. |  | 0.010059 | 0.008165 | 0.014489 | 0.011034 | 0.029269 | 0.039992 | 0.0040825 | 0.0043894 | 0.020913 | 0.02137 | 0.018619 | 0.0063246 | 0.0075277 | 0.0083666 | 0.032496 | 0.0098319 | 0.0075277 | 0.0075277 | 0.005164 | 0.019408 | 0.0083666 | 0.0098319 | 0.018619 | 0.033507 | 0.010602 |  | 0.026065 | 0.010488
| Rel.Std.Dev. |  | 1.793% | 1.503% | 2.481% | 1.826% | 4.191% | 5.312% | 0.69% | 1.009% | 3.486% | 3.674% | 2.971% | 1.129% | 1.317% | 1.455% | 6.118% | 1.73% | 1.348% | 1.294% | 0.9995% | 3.539% | 1.564% | 1.849% | 3.558% | 5.546% | 1.77% |  | 4.353% | 1.734%
| PDM3 |  | -1.888% | -4.985% | 2.111% | 5.682% | 22.12% | 31.66% | 3.467% | -23.94% | 4.902% | 1.719% | 9.588% | -2.07% | -0.03021% | 0.5527% | -7.118% | -0.6131% | -2.362% | 1.719% | -9.648% | -4.111% | -6.442% | -7.025% | -8.482% | 5.659% | 4.75% |  | 4.721% | 5.799%
| Z-Score (Absolute) |  | 0.45 | 0.91 | 0.15 | 0.68 | 3.12 | 4.54 | 0.35 | 3.73 | 0.56 | 0.09 | 1.26 | 0.48 | 0.17 | 0.09 | 1.23 | 0.26 | 0.52 | 0.09 | 1.6 | 0.78 | 1.13 | 1.21 | 1.43 | 0.67 | 0.54 | 10.96 | 0.53 | 0.69
| Analytical results for Cd in OREAS 284 (Certified Value 0.084 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | < 0.2 | 0.1 | 0.080499 | < 1 | 0.11 | 0.6 | 0.04 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.074341 | < 0.1 | 0.08 | 0.09 | 0.09 | 0.08 | 0.09 | < 0.5 | 0.1 | 0.11 | 0.531 | 0.093 | 0.09 |  |  |  |  | 1
| 1 | 2 | < 0.2 | 0.08 | 0.073707 | < 1 | 0.11 | 0.6 | 0.09 | 0.08 | 0.09 | 0.09 | 0.07 | 0.07 | 0.085507 | < 0.1 | 0.09 | 0.08 | 0.09 | 0.08 | 0.08 | < 0.5 | 0.08 | 0.11 | 0.531 | 0.09 | 0.06 |  |  |  |  | 27
| 1 | 3 | < 0.2 | 0.1 | 0.091988 | < 1 | 0.07 | 0.6 | 0.09 | 0.09 | 0.08 | 0.1 | 0.09 | 0.09 | 0.087314 | < 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.08 | < 0.5 | 0.09 | 0.11 | 0.594 | 0.08 | 0.1 |  |  |  |  | 16
| 1 | 4 | < 0.2 | 0.09 | 0.081996 | < 1 | 0.08 | 0.6 | 0.08 | 0.07 | 0.1 | 0.09 | 0.08 | 0.07 | 0.073469 | < 0.1 | 0.09 | 0.1 | 0.08 | 0.09 | 0.09 | < 0.5 | 0.09 | 0.1 | 0.459 | 0.08 | 0.06 |  |  |  |  | 0.083833
| 1 | 5 | < 0.2 | 0.08 | 0.075296 | < 1 | 0.11 | 0.6 | 0.07 | 0.09 | 0.08 | 0.09 | 0.08 | 0.08 | 0.085216 | < 0.1 | 0.08 | 0.09 | 0.09 | 0.08 | 0.09 | < 0.5 | 0.1 | 0.1 | 0.594 | 0.08 | 0.07 |  |  |  |  | 80
| 1 | 6 | < 0.2 | 0.09 | 0.084348 | < 1 | 0.11 | 0.6 | 0.08 | 0.06 | 0.09 | 0.07 | 0.08 | 0.06 | 0.094302 | < 0.1 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | < 0.5 | 0.08 | 0.1 | 0.504 | 0.088 | 0.07
| Mean |  |  | 0.09 | 0.081306 |  | 0.098333 | 0.6 | 0.075 | 0.078333 | 0.086667 | 0.086667 | 0.08 | 0.075 | 0.083358 |  | 0.086667 | 0.088333 | 0.086667 | 0.083333 | 0.086667 |  | 0.09 | 0.105 | 0.5355 | 0.085167 | 0.075
| Median |  |  | 0.09 | 0.081248 |  | 0.11 | 0.6 | 0.08 | 0.08 | 0.085 | 0.09 | 0.08 | 0.075 | 0.085361 |  | 0.09 | 0.09 | 0.09 | 0.08 | 0.09 |  | 0.09 | 0.105 | 0.531 | 0.084 | 0.07
| Std Dev. |  |  | 0.0089443 | 0.0066085 |  | 0.018348 | 0 | 0.018708 | 0.01169 | 0.008165 | 0.010328 | 0.0063246 | 0.010488 | 0.0080321 |  | 0.005164 | 0.0075277 | 0.005164 | 0.005164 | 0.005164 |  | 0.0089443 | 0.0054772 | 0.052401 | 0.0058793 | 0.016432
| Rel.Std.Dev. |  |  | 9.938% | 8.128% |  | 18.66% | 0% | 24.94% | 14.92% | 9.421% | 11.92% | 7.906% | 13.98% | 9.636% |  | 5.958% | 8.522% | 5.958% | 6.197% | 5.958% |  | 9.938% | 5.216% | 9.785% | 6.903% | 21.91%
| PDM3 |  |  | 7.356% | -3.015% |  | 17.3% | 615.7% | -10.54% | -6.56% | 3.38% | 3.38% | -4.572% | -10.54% | -0.5666% |  | 3.38% | 5.368% | 3.38% | -0.5962% | 3.38% |  | 7.356% | 25.25% | 538.8% | 1.591% | -10.54%
| Z-Score (Absolute) |  | 1.35 | 0.34 | 0.54 | 41.81 | 1.18 | 51.92 | 1.18 | 0.84 | 0 | 0 | 0.67 | 1.18 | 0.33 | 3.71 | 0 | 0.17 | 0 | 0.34 | 0 | 16.52 | 0.34 | 1.85 | 45.4 | 0.15 | 1.18
| Analytical results for Ce in OREAS 284 (Certified Value 40.4 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | [9 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [9 empty] | 1
| 1 | 1 | 44.2046 | 39.067 | 42 | 49.1111 | 38.6 | 42.4 | 40.2 | 37.5 | 34.4 | 39.8453 | 37 | 34 | 56.7 | 33.1 | 41.49 | 45.659 | 16 | 86.64 | 47.7 | 41.68 | [9 empty] | 1
| 1 | 2 | 43.5312 | 37.548 | 40.3 | 49.7985 | 38.2 | 41.9 | 38.2 | 35.9 | 36.2 | 40.2725 | 39 | 34.49 | 57.9 | 32.6 | 41.22 | 49.685 | 14.9 | 87.36 | 45.1 | 43.07 | [9 empty] | 28
| 1 | 3 | 42.8007 | 36.574 | 40.2 | 48.0308 | 40.7 | 38.9 | 38.4 | 37.2 | 35 | 40.7601 | 39 | 34.61 | 58.8 | 33.2 | 42.23 | 51.218 | 16.5 | 84.08 | 43.2 | 43.1 | [9 empty] | 16
| 1 | 4 | 42.0306 | 35.347 | 39.5 | 46.063 | 40.2 | 44.4 | 36.1 | 39.2 | 33.8 | 41.3959 | 40 | 34.21 | 57.9 | 32.8 | 41.99 | 48.298 | 16 | 86.64 | 45.9 | 44.67 | [9 empty] | 40.4456
| 1 | 5 | 44.261 | 35.968 | 42 | 51.257 | 39.5 | 42.6 | 40 | 38.1 | 35.8 | 39.7089 | 41 | 33.46 | 57.4 | 32.8 | 45.27 | 49.082 | 15.4 | 84.32 | 39.2 | 43.61 | [9 empty] | 81
| 1 | 6 | 44.913 | 31.598 | 40.8 | 49.1014 | 34.8 | 40 | 39.8 | 36.9 | 35 | 40.8215 | 41 | 34.5 | 57.5 | 33.3 | 45.35 | 50.658 | 14.7 | 85.6 | 42.6 | 44.69
| Mean |  | 43.6235 | 36.017 | 40.8 | 48.8936 | 38.6667 | 41.7 | 38.7833 | 37.4667 | 35.0333 | 40.4674 | 39.5 | 34.2117 | 57.7 | 32.9667 | 42.925 | 49.1 | 15.5833 | 85.7733 | 43.95 | 43.47
| Median |  | 43.8679 | 36.271 | 40.55 | 49.1063 | 39.05 | 42.15 | 39.1 | 37.35 | 35 | 40.5163 | 39.5 | 34.35 | 57.7 | 32.95 | 42.11 | 49.3835 | 15.7 | 86.12 | 44.15 | 43.355
| Std Dev. |  | 1.0611 | 2.5273 | 1.0178 | 1.7465 | 2.1144 | 1.9677 | 1.5626 | 1.1183 | 0.88015 | 0.64398 | 1.5166 | 0.43106 | 0.6957 | 0.27325 | 1.8817 | 1.9869 | 0.70261 | 1.3438 | 2.971 | 1.1357
| Rel.Std.Dev. |  | 2.432% | 7.017% | 2.495% | 3.572% | 5.468% | 4.719% | 4.029% | 2.985% | 2.512% | 1.591% | 3.839% | 1.26% | 1.206% | 0.8289% | 4.384% | 4.047% | 4.509% | 1.567% | 6.76% | 2.613%
| PDM3 |  | 7.857% | -10.95% | 0.8763% | 20.89% | -4.398% | 3.102% | -4.11% | -7.365% | -13.38% | 0.05383% | -2.338% | -15.41% | 42.66% | -18.49% | 6.13% | 21.4% | -61.47% | 112.1% | 8.665% | 7.478%
| Z-Score (Absolute) |  | 0.62 | 0.96 | 0.03 | 1.72 | 0.41 | 0.22 | 0.38 | 0.66 | 1.16 | 0.03 | 0.24 | 1.34 | 3.55 | 1.59 | 0.48 | 1.76 | 5.21 | 9.39 | 0.69 | 0.59
| Analytical results for Co in OREAS 284 (Certified Value 11.9 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 1
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 2
| 1 | 1 | 15.46 | 12.2979 | 12.5 | 17.1309 | 12.6724 | 13.4 | 12.0461 | 11 | 10.2 | 11.8 | 13 | 11.6 | 11.6 | 12 | 10.8735 | 11.3 | 11.5 | 11.7 | 11.7 | 11.6 | 12.5 | 12 | 11.4 | 12.34 | 9.5 | 12.25 | 12.6 | 11.5 |  | 1
| 1 | 2 | 15.88 | 11.9329 | 12.8 | 16.7432 | 12.5735 | 12.8 | 12.4924 | 11 | 10.6 | 11.6 | 13 | 11.4 | 12 | 12.2 | 11.7675 | 11.2 | 11.2 | 11.3 | 11.6 | 12.1 | 11.6 | 12 | 12 | 12.49 | 9.1 | 11.99 | 12.3 | 12 |  | 16
| 1 | 3 | 15.71 | 11.9574 | 12.6 | 16.5771 | 12.4665 | 12.7 | 10.1394 | 11 | 11.4 | 11.9 | 11.6 | 11.4 | 11.5 | 11.4 | 11.1895 | 11.2 | 11.9 | 11.8 | 11.6 | 11.7 | 11.9 | 12 | 12.2 | 12.89 | 9.09 | 11.89 | 12.2 | 12.1 |  | 16
| 1 | 4 | 15.93 | 12.1532 | 12.2 | 16.2111 | 12.4395 | 12.8 | 10.5005 | 11 | 11.3 | 11.8 | 13.3 | 11.2 | 12.1 | 11.2 | 11.9258 | 11 | 11.5 | 11.6 | 11.6 | 12.3 | 12 | 11 | 12 | 12.9 | 9.21 | 11.18 | 12.6 | 11.9 |  | 11.914
| 1 | 5 | 16.05 | 11.9259 | 12.5 | 16.7624 | 12.3564 | 13 | 10.564 | 11 | 11.3 | 11.8 | 13.4 | 12 | 12 | 11.8 | 11.3631 | 11.1 | 11.3 | 11.4 | 11.7 | 12.3 | 12.3 | 11 | 12.4 | 12.6 | 9.1 | 11.96 | 11.7 | 11.7 |  | 82
| 1 | 6 | 15.73 | 12.2026 | 12.5 | 17.5467 | 12.5516 | 12.3 | 10.4999 | 11 | 11 | 12.2 | 12.5 | 11.7 | 11.9 | 11.7 | 12.1752 | 11.2 | 11.8 | 11.3 | 11.5 | 12 | 12.1 | 11 | 11.6 | 12.39 | 8.85 | 11.77 | 12.2 | 12.3
| Mean |  | 15.7933 | 12.0783 | 12.5167 | 16.8286 | 12.51 | 12.8333 | 11.0404 | 11 | 10.9667 | 11.85 | 12.8 | 11.55 | 11.85 | 11.7167 | 11.5491 | 11.1667 | 11.5333 | 11.5167 | 11.6167 | 12 | 12.0667 | 11.5 | 11.9333 | 12.6017 | 9.1417 | 11.84 | 12.2667 | 11.9167
| Median |  | 15.805 | 12.0553 | 12.5 | 16.7528 | 12.5091 | 12.8 | 10.5322 | 11 | 11.15 | 11.8 | 13 | 11.5 | 11.95 | 11.75 | 11.5653 | 11.2 | 11.5 | 11.5 | 11.6 | 12.05 | 12.05 | 11.5 | 12 | 12.545 | 9.1 | 11.925 | 12.25 | 11.95
| Std Dev. |  | 0.20695 | 0.16018 | 0.19408 | 0.46115 | 0.11187 | 0.36148 | 0.97387 | 0 | 0.4761 | 0.19748 | 0.66633 | 0.28107 | 0.2429 | 0.37103 | 0.49032 | 0.10328 | 0.27325 | 0.2137 | 0.075277 | 0.29665 | 0.31411 | 0.54772 | 0.37238 | 0.24408 | 0.21179 | 0.36 | 0.33267 | 0.28577
| Rel.Std.Dev. |  | 1.31% | 1.326% | 1.551% | 2.74% | 0.8942% | 2.817% | 8.821% | 0% | 4.341% | 1.667% | 5.206% | 2.434% | 2.05% | 3.167% | 4.246% | 0.9249% | 2.369% | 1.856% | 0.648% | 2.472% | 2.603% | 4.763% | 3.121% | 1.937% | 2.317% | 3.041% | 2.712% | 2.398%
| PDM3 |  | 32.56% | 1.38% | 5.059% | 41.25% | 5.003% | 7.717% | -7.333% | -7.671% | -7.951% | -0.537% | 7.437% | -3.055% | -0.537% | -1.656% | -3.063% | -6.273% | -3.195% | -3.335% | -2.495% | 0.722% | 1.282% | -3.475% | 0.1625% | 5.772% | -23.27% | -0.6209% | 2.96% | 0.02257%
| Z-Score (Absolute) |  | 7.36 | 0.37 | 1.19 | 9.31 | 1.18 | 1.79 | 1.59 | NA | 1.72 | 0.06 | 1.72 | 0.63 | 0.06 | 0.31 | 0.63 | 1.35 | 0.66 | 0.69 | 0.5 | 0.22 | 0.34 | NA | 0.09 | 1.35 | 5.16 | 0.08 | 0.72 | 0.06
Results from laboratories 1.11 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 284 (Certified Value 52 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  | 0
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 1
| 1 | 1 | 48.95 | 53.3924 | 53 | 56.337 | 53.689 | 52 | 52.9221 | 48 | 47.4 | 53 | 55 | 54 | 52 | 52 | 49.7027 | 56 | 55 | 52 | 55.1 | 51.1 | 53 | 51 | 51 | 60.42 | 59 | 47.184 | 25.3 | 52 |  | 1
| 1 | 2 | 50.4 | 50.6327 | 52 | 55.7087 | 51.3252 | 51 | 55.4922 | 48 | 49.9 | 54 | 57 | 54 | 52 | 51 | 48.8088 | 57 | 56 | 51 | 55.7 | 50 | 51 | 50 | 54 |  | 58 | 45.752 | 24 | 52 |  | 17
| 1 | 3 | 49.74 | 50.7751 | 52 | 54.3025 | 51.7068 | 50 | 49.2355 | 47 | 49.6 | 54 | 53 | 53 | 52 | 51 | 51.1124 | 55 | 55 | 52 | 54.6 | 48.9 | 51 | 51 | 56 | 61.14 | 58 | 44.632 | 24 | 55 |  | 16
| 1 | 4 | 50.3 | 52.0607 | 53 | 52.3915 | 52.0146 | 50 | 48.4601 | 47 | 50 | 53 | 53 | 52 | 53 | 51 | 51.3406 | 54 | 55 | 52 | 54.8 | 51.1 | 52 | 49 | 54 | 59.74 | 59 | 42.384 | 25 | 53 |  | 52.1665
| 1 | 5 | 51.14 | 51.4397 | 52 | 54.558 | 51.9725 | 52 | 49.8156 | 47 | 49.8 | 54 | 54 | 54 | 53 | 52 | 53.4532 | 54 | 56 | 52 | 55.5 | 52.3 | 52 | 51 | 57 | 58.82 | 59 | 44.832 | 23 | 53 |  | 83
| 1 | 6 | 49.76 | 52.8452 | 53 | 59.0184 | 52.2262 | 47 | 49.7982 | 47 | 47.3 | 50 | 50 | 53 | 52 | 53 | 53.3304 | 61 | 56 | 51 | 54.3 | 50.9 | 52 | 49 | 53 | 58.2 | 58 | 44.008 | 23.9 | 56
| Mean |  | 50.0483 | 51.8576 | 52.5 | 55.386 | 52.1557 | 50.3333 | 50.9539 | 47.3333 | 49 | 53 | 53.6667 | 53.3333 | 52.3333 | 51.6667 | 51.2913 | 56.1667 | 55.5 | 51.6667 | 55 | 50.7167 | 51.8333 | 50.1667 | 54.1667 | 59.664 | 58.5 | 44.7987 | 24.2 | 53.5
| Median |  | 50.03 | 51.7502 | 52.5 | 55.1333 | 51.9936 | 50.5 | 49.8069 | 47 | 49.7 | 53.5 | 53.5 | 53.5 | 52 | 51.5 | 51.2265 | 55.5 | 55.5 | 52 | 54.95 | 51 | 52 | 50.5 | 54 | 59.74 | 58.5 | 44.732 | 24 | 53
| Std Dev. |  | 0.74341 | 1.1153 | 0.54772 | 2.2374 | 0.81251 | 1.8619 | 2.6934 | 0.5164 | 1.2853 | 1.5492 | 2.3381 | 0.8165 | 0.5164 | 0.8165 | 1.8736 | 2.6394 | 0.54772 | 0.5164 | 0.53666 | 1.1531 | 0.75277 | 0.98319 | 2.137 | 1.1844 | 0.54772 | 1.6176 | 0.83187 | 1.6432
| Rel.Std.Dev. |  | 1.485% | 2.151% | 1.043% | 4.04% | 1.558% | 3.699% | 5.286% | 1.091% | 2.623% | 2.923% | 4.357% | 1.531% | 0.9867% | 1.58% | 3.653% | 4.699% | 0.9869% | 0.9995% | 0.9757% | 2.274% | 1.452% | 1.96% | 3.945% | 1.985% | 0.9363% | 3.611% | 3.437% | 3.071%
| PDM3 |  | -4.06% | -0.592% | 0.6394% | 6.172% | -0.02061% | -3.514% | -2.324% | -9.265% | -6.07% | 1.598% | 2.876% | 2.237% | 0.3199% | -0.9581% | -1.678% | 7.668% | 6.39% | -0.9581% | 5.432% | -2.779% | -0.6386% | -3.834% | 3.834% | 14.37% | 12.14% | -14.12% | -53.61% | 2.556%
| Z-Score (Absolute) |  | 0.79 | 0.06 | 0.2 | 1.37 | 0.06 | 0.68 | 0.43 | 1.89 | 1.22 | 0.4 | 0.67 | 0.54 | 0.13 | 0.14 | 0.29 | 1.68 | 1.41 | 0.14 | 1.21 | 0.52 | 0.07 | 0.74 | 0.87 | 3.1 | 2.63 | 2.92 | 11.25 | 0.6
| Analytical results for Cs in OREAS 284 (Certified Value 9.48 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [11 empty] | 3
| 1 | 1 | 9.3554 | 9.58 | 8.65 | 9.39 | 10.75 | 9.58 | 9.85 | 8.96 | 9.09 | 9.3 | 9.08 | 9.94 | 9.24 | 9.34 | 8.97 | 7.74 | 10.8 | 9.4 | [11 empty] | 1
| 1 | 2 | 9.5852 | 9.47 | 8.4 | 9.24 | 10.7 | 9.46 | 9.84 | 9.41 | 9.248 | 9.5 | 9.65 | 10.2 | 9.19 | 9.03 | 9.7 | 7.45 | 10.6 | 9.73 | [11 empty] | 31
| 1 | 3 | 9.255 | 9.46 | 8.59 | 9.63 | 9.89 | 9.41 | 9.41 | 8.86 | 9.036 | 9.5 | 9.31 | 10.1 | 9.47 | 9.22 | 9.84 | 7.45 | 10.8 | 9.75 | [11 empty] | 16
| 1 | 4 | 9.3425 | 9.28 | 8.4 | 9.61 | 11.35 | 9.21 | 10.1 | 8.78 | 9.4261 | 9.5 | 9.38 | 10 | 9.38 | 9.16 | 9.53 | 7.41 | 10.8 | 9.69 | [11 empty] | 9.4776
| 1 | 5 | 9.6525 | 9.3 | 8.67 | 9.34 | 10.5 | 9.87 | 9.84 | 9.2 | 9.0957 | 9.6 | 9.31 | 9.94 | 9.37 | 9.45 | 9.94 | 7.32 | 10.2 | 9.65 | [11 empty] | 84
| 1 | 6 | 9.5485 | 9.54 | 8.35 | 9.6 | 9.42 | 9.56 | 9.68 | 9.01 | 9.4269 | 9.7 | 9.51 | 9.92 | 9.48 | 9.47 | 9.5 | 7.1 | 10.5 | 10.6
| Mean |  | 9.4565 | 9.4383 | 8.51 | 9.4683 | 10.435 | 9.515 | 9.7867 | 9.0367 | 9.2204 | 9.5167 | 9.3733 | 10.0167 | 9.355 | 9.2783 | 9.58 | 7.4117 | 10.6167 | 9.8033
| Median |  | 9.4519 | 9.465 | 8.495 | 9.495 | 10.6 | 9.51 | 9.84 | 8.985 | 9.1718 | 9.5 | 9.345 | 9.97 | 9.375 | 9.28 | 9.615 | 7.43 | 10.7 | 9.71
| Std Dev. |  | 0.15954 | 0.12336 | 0.14241 | 0.1663 | 0.68383 | 0.21879 | 0.22853 | 0.23226 | 0.1745 | 0.13292 | 0.19459 | 0.1113 | 0.11845 | 0.17268 | 0.34427 | 0.20798 | 0.24014 | 0.41035
| Rel.Std.Dev. |  | 1.687% | 1.307% | 1.673% | 1.756% | 6.553% | 2.299% | 2.335% | 2.57% | 1.893% | 1.397% | 2.076% | 1.111% | 1.266% | 1.861% | 3.594% | 2.806% | 2.262% | 4.186%
| PDM3 |  | -0.2223% | -0.414% | -10.21% | -0.09743% | 10.1% | 0.395% | 3.261% | -4.652% | -2.713% | 0.4125% | -1.1% | 5.688% | -1.293% | -2.102% | 1.081% | -21.8% | 12.02% | 3.437%
| Z-Score (Absolute) |  | 0.02 | 0.08 | 3.01 | 0.02 | 3.08 | 0.17 | 1.03 | 1.35 | 0.77 | 0.17 | 0.28 | 1.75 | 0.34 | 0.58 | 0.37 | 6.49 | 3.65 | 1.08
| Analytical results for Cu in OREAS 284 (Certified Value 27.4 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 2
| 1 | 1 | 28.136 | 29.6 | 31.2977 | 31.9972 | 31.1 | 27.534 | 29 | 31 | 26.44 | 26.7 | 28.2 | 27.8 | 26 | 27.6 | 25.9653 | 21.4 | 27.6 | 25 | 27.5 | 25.87 | 28.1 | 36 | 28.4 |  | 25.5 | 25.76 | 33.3 | 28.1 |  | 1
| 1 | 2 | 27.4866 | 29.8 | 28.6507 | 30.3415 | 29.1 | 28.9396 | 29 | 30 | 26.72 | 27.4 | 28.2 | 26.2 | 26.8 | 27.1 | 25.7209 | 21.4 | 28.6 | 25 | 27.1 | 26.48 | 28.6 | 28 | 28 | 26.05 | 24.9 | 26.38 | 32.7 | 27.9 |  | 18
| 1 | 3 | 27.2416 | 29.9 | 26.1964 | 30.3336 | 29.9 | 25.1257 | 29 | 31 | 27.56 | 27.2 | 25.7 | 26.7 | 25.7 | 26 | 27.0839 | 20.9 | 27.2 | 25 | 27.1 | 25.74 | 28.3 | 28 | 28.6 | 26.26 | 24.7 | 25.81 | 32.7 | 28.6 |  | 16
| 1 | 4 | 27.7001 | 29.3 | 25.4672 | 33.2963 | 31 | 25.3061 | 28 | 30 | 27.76 | 27.6 | 29.8 | 26.5 | 27.1 | 25.5 | 27.5579 | 20.3 | 29 | 25 | 26.3 | 26.27 | 28.6 | 27 | 27.9 | 26.32 | 25 | 24.85 | 33.3 | 29.2 |  | 27.4121
| 1 | 5 | 27.1914 | 29.4 | 27.6393 | 31.9421 | 30.6 | 24.3736 | 29 | 30 | 26.46 | 26.1 | 28.8 | 27.9 | 26 | 26.2 | 27.8958 | 20.5 | 27.3 | 25 | 26.8 | 26.81 | 28.4 | 27 | 28.9 | 26.31 | 24.5 | 27.23 | 31.7 | 29.3 |  | 85
| 1 | 6 | 27.4714 | 29.2 | 29.5361 | 31.555 | 30.1 | 26.6263 | 29 | 28 | 27.1 | 27.1 | 26 | 27.1 | 26.1 | 26.1 | 27.4002 | 20.2 | 28.5 | 25 | 26.6 | 26.9 | 28.4 | 26 | 28 | 25.45 | 24.4 | 27.02 | 32.8 | 29.4
| Mean |  | 27.5379 | 29.5333 | 28.1312 | 31.5776 | 30.3 | 26.3176 | 28.8333 | 30 | 27.0067 | 27.0167 | 27.7833 | 27.0333 | 26.2833 | 26.4167 | 26.9373 | 20.7833 | 28.0333 | 25 | 26.9 | 26.345 | 28.4 | 28.6667 | 28.3 | 26.078 | 24.8333 | 26.175 | 32.75 | 28.75
| Median |  | 27.479 | 29.5 | 28.145 | 31.7486 | 30.35 | 25.9662 | 29 | 30 | 26.91 | 27.15 | 28.2 | 26.9 | 26.05 | 26.15 | 27.2421 | 20.7 | 28.05 | 25 | 26.95 | 26.375 | 28.4 | 27.5 | 28.2 | 26.26 | 24.8 | 26.095 | 32.75 | 28.9
| Std Dev. |  | 0.34611 | 0.28048 | 2.1611 | 1.126 | 0.75631 | 1.7131 | 0.40825 | 1.0954 | 0.56294 | 0.54191 | 1.6105 | 0.69761 | 0.54191 | 0.77825 | 0.89042 | 0.53448 | 0.7607 | 0 | 0.42426 | 0.47744 | 0.18974 | 3.6697 | 0.4 | 0.36765 | 0.39833 | 0.88658 | 0.58566 | 0.6473
| Rel.Std.Dev. |  | 1.257% | 0.9497% | 7.682% | 3.566% | 2.496% | 6.509% | 1.416% | 3.651% | 2.084% | 2.006% | 5.797% | 2.581% | 2.062% | 2.946% | 3.306% | 2.572% | 2.714% | 0% | 1.577% | 1.812% | 0.6681% | 12.8% | 1.413% | 1.41% | 1.604% | 3.387% | 1.788% | 2.251%
| PDM3 |  | 0.4588% | 7.738% | 2.623% | 15.2% | 10.54% | -3.993% | 5.185% | 9.441% | -1.479% | -1.443% | 1.354% | -1.382% | -4.118% | -3.631% | -1.732% | -24.18% | 2.266% | -8.799% | -1.868% | -3.893% | 3.604% | 4.577% | 3.239% | -4.867% | -9.407% | -4.513% | 19.47% | 4.881%
| Z-Score (Absolute) |  | 0.16 | 1.43 | 0.54 | 2.72 | 1.91 | 0.61 | 0.98 | 1.72 | 0.18 | 0.17 | 0.32 | 0.16 | 0.64 | 0.55 | 0.22 | 4.13 | 0.47 | 1.45 | 0.24 | 0.6 | 0.71 | 0.88 | 0.64 | 0.77 | 1.56 | 0.7 | 3.47 | 0.93
| Analytical results for Dy in OREAS 284 (Certified Value 2.81 ppm) | [30 empty] | Certified
Dy | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.29 |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*MS |  | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.5g |  | 3
| 1 | 1 | 1.03: 2.9001; 1.04: 2.73; 1.05: 2.9; 1.07: 3.0156; 1.09: 2.5 |  | 1
| 1 | 2 | 1.03: 2.9991; 1.04: 2.701; 1.05: 2.8; 1.07: 3.1617; 1.09: 2.35 |  | 33
| 1 | 3 | 1.03: 2.8863; 1.04: 2.645; 1.05: 3; 1.07: 2.9995; 1.09: 2.37 |  | 16
| 1 | 4 | 1.03: 2.9119; 1.04: 2.612; 1.05: 2.9; 1.07: 3.0331; 1.09: 2.42 |  | 2.8114
| 1 | 5 | 1.03: 3.0407; 1.04: 2.678; 1.05: 3; 1.07: 3.2536; 1.09: 2.39 |  | 86
| 1 | 6 | 1.03: 2.9826; 1.04: 2.75; 1.05: 2.9; 1.07: 3.2511; 1.09: 2.26
| Mean |  | 1.03: 2.9535; 1.04: 2.686; 1.05: 2.9167; 1.07: 3.1191; 1.09: 2.3817
| Median |  | 1.03: 2.9473; 1.04: 2.6895; 1.05: 2.9; 1.07: 3.0974; 1.09: 2.38
| Std Dev. |  | 1.03: 0.062642; 1.04: 0.051919; 1.05: 0.075277; 1.07: 0.1181; 1.09: 0.079352
| Rel.Std.Dev. |  | 1.03: 2.121%; 1.04: 1.933%; 1.05: 2.581%; 1.07: 3.786%; 1.09: 3.332%
| PDM3 |  | 1.03: 5.054%; 1.04: -4.46%; 1.05: 3.745%; 1.07: 10.95%; 1.09: -15.28%
| Z-Score (Absolute) |  | 1.03: 0.12; 1.04: 0.77; 1.05: 0; 1.07: 0.67; 1.09: 1.78
| Analytical results for Er in OREAS 284 (Certified Value 1 ppm) | [30 empty] | Certified
Er | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.29 |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*MS |  | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.5g |  | 2
| 1 | 1 | 1.03: 1.0498; 1.04: 0.99; 1.05: 1; 1.07: 1.1134; 1.09: 0.9 |  | 1
| 1 | 2 | 1.03: 1.0677; 1.04: 0.967; 1.05: 1; 1.07: 1.1658; 1.09: 0.84 |  | 34
| 1 | 3 | 1.03: 1.0424; 1.04: 0.957; 1.05: 1; 1.07: 1.1092; 1.09: 0.84 |  | 16
| 1 | 4 | 1.03: 1.036; 1.04: 0.945; 1.05: 1; 1.07: 1.1377; 1.09: 0.84 |  | 0.99932
| 1 | 5 | 1.03: 1.0771; 1.04: 0.948; 1.05: 1; 1.07: 1.1928; 1.09: 0.84 |  | 87
| 1 | 6 | 1.03: 1.0745; 1.04: 0.889; 1.05: 1; 1.07: 1.2371; 1.09: 0.79
| Mean |  | 1.03: 1.0579; 1.04: 0.94933; 1.05: 1; 1.07: 1.1593; 1.09: 0.84167
| Median |  | 1.03: 1.0587; 1.04: 0.9525; 1.05: 1; 1.07: 1.1518; 1.09: 0.84
| Std Dev. |  | 1.03: 0.0175; 1.04: 0.03372; 1.05: 0; 1.07: 0.0496; 1.09: 0.034881
| Rel.Std.Dev. |  | 1.03: 1.654%; 1.04: 3.552%; 1.05: 0%; 1.07: 4.278%; 1.09: 4.144%
| PDM3 |  | 1.03: 5.864%; 1.04: -5.002%; 1.05: 0.0684%; 1.07: 16.01%; 1.09: -15.78%
| Z-Score (Absolute) |  | 1.03: 0.67; 1.04: 0.59; 1.05: 0; 1.07: 1.86; 1.09: 1.84
| Analytical results for Eu in OREAS 284 (Certified Value 0.39 ppm) | [30 empty] | Certified
Eu | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.29 |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*MS |  | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.5g |  | 2
| 1 | 1 | 1.03: 0.39251; 1.04: 0.312; 1.05: 0.4; 1.07: 0.4781; 1.09: 0.34 |  | 1
| 1 | 2 | 1.03: 0.37874; 1.04: 0.311; 1.05: 0.4; 1.07: 0.5223; 1.09: 0.32 |  | 35
| 1 | 3 | 1.03: 0.38196; 1.04: 0.312; 1.05: 0.4; 1.07: 0.5049; 1.09: 0.33 |  | 16
| 1 | 4 | 1.03: 0.38308; 1.04: 0.308; 1.05: 0.4; 1.07: 0.495; 1.09: 0.33 |  | 0.38513
| 1 | 5 | 1.03: 0.40239; 1.04: 0.309; 1.05: 0.4; 1.07: 0.523; 1.09: 0.31 |  | 88
| 1 | 6 | 1.03: 0.3825; 1.04: 0.279; 1.05: 0.4; 1.07: 0.5483; 1.09: 0.3
| Mean |  | 1.03: 0.38686; 1.04: 0.30517; 1.05: 0.4; 1.07: 0.51193; 1.09: 0.32167
| Median |  | 1.03: 0.38279; 1.04: 0.31; 1.05: 0.4; 1.07: 0.5136; 1.09: 0.325
| Std Dev. |  | 1.03: 0.0089012; 1.04: 0.012922; 1.05: 6.0809e-17; 1.07: 0.024631; 1.09: 0.01472
| Rel.Std.Dev. |  | 1.03: 2.301%; 1.04: 4.234%; 1.05: 1.52e-14%; 1.07: 4.811%; 1.09: 4.576%
| PDM3 |  | 1.03: 0.4513%; 1.04: -20.76%; 1.05: 3.862%; 1.07: 32.93%; 1.09: -16.48%
| Z-Score (Absolute) |  | 1.03: 0; 1.04: 0.84; 1.05: 0.14; 1.07: 1.29; 1.09: 0.67
| Analytical results for Fe in OREAS 284 (Certified Value 3.08 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 3.1022 | 3.14 | 3.1786 | 3.3745 | 3.51 | 3.0775 | 3.18 | 1.6044 | 2.76 | 3.13 | 3.34 | 3.04 | 3.09 | 2.96 | 3.1361 | 2.75 | 2.86 | 3.05 | 2.85 | 2.86 | 3.13 | 3.22 | 2.87 | 3.142 | 3.029 | 3.447 | 3.24 | 3.11 |  | 1
| 1 | 2 | 3.0037 | 3.11 | 3.2262 | 3.1708 | 3.34 | 3.2701 | 3.2 | 1.5927 | 2.92 | 3.16 | 3.34 | 3.06 | 3.05 | 2.96 | 3.0852 | 2.74 | 2.89 | 2.98 | 2.9 | 2.95 | 3.02 | 3.17 | 3.03 | 3.101 | 2.938 | 3.596 | 3.12 | 3.18 |  | 11
| 1 | 3 | 2.973 | 3.1 | 3.1348 | 3.2302 | 3.37 | 2.8973 | 3.17 | 1.6311 | 2.93 | 3.14 | 3.25 | 3.03 | 3.08 | 2.96 | 3.0898 | 2.75 | 2.83 | 3.02 | 2.83 | 2.87 | 3.09 | 3.24 | 3.11 | 3.305 | 2.945 | 3.297 | 3.03 | 3.2 |  | 16
| 1 | 4 | 3.047 | 3.17 | 3.0665 | 3.3427 | 3.44 | 2.8682 | 3.14 | 1.5809 | 2.94 | 3.14 | 3.23 | 2.98 | 3.11 | 2.96 | 3.1373 | 2.78 | 2.88 | 2.99 | 2.82 | 2.95 | 3.14 | 3.12 | 3.01 | 3.105 | 3.015 | 3.186 | 3.18 | 3.21 |  | 3.0825
| 1 | 5 | 3.0515 | 3.07 | 3.2364 | 3.2705 | 3.42 | 2.922 | 3.14 | 1.5792 | 2.94 | 3.18 | 3.26 | 3.07 | 3.11 | 3.03 | 3.1524 | 2.81 | 2.91 | 2.91 | 2.87 | 3.01 | 3.11 | 3.18 | 3.21 | 3.435 | 2.966 | 3.317 | 2.9 | 3.22 |  | 89
| 1 | 6 | 3.1176 | 3.15 | 3.0969 | 3.2465 | 3.25 | 2.9868 | 3.14 | 1.6123 | 2.81 | 2.94 | 3.11 | 3.04 | 3.08 | 3.1 | 3.1307 | 2.82 | 2.88 | 2.98 | 2.79 | 2.94 | 3.12 | 3.05 | 2.97 | 3.292 | 2.98 | 3.382 | 3.01 | 3.28
| Mean |  | 3.0492 | 3.1233 | 3.1565 | 3.2725 | 3.3883 | 3.0036 | 3.1617 | 1.6001 | 2.8833 | 3.115 | 3.255 | 3.0367 | 3.0867 | 2.995 | 3.1219 | 2.775 | 2.875 | 2.9883 | 2.8433 | 2.93 | 3.1017 | 3.1633 | 3.0333 | 3.23 | 2.9788 | 3.3708 | 3.08 | 3.2
| Median |  | 3.0492 | 3.125 | 3.1567 | 3.2585 | 3.395 | 2.9544 | 3.155 | 1.5985 | 2.925 | 3.14 | 3.255 | 3.04 | 3.085 | 2.96 | 3.1334 | 2.765 | 2.88 | 2.985 | 2.84 | 2.945 | 3.115 | 3.175 | 3.02 | 3.217 | 2.973 | 3.3495 | 3.075 | 3.205
| Std Dev. |  | 0.055425 | 0.036697 | 0.069044 | 0.075029 | 0.089759 | 0.15037 | 0.025626 | 0.019943 | 0.078145 | 0.087579 | 0.085029 | 0.031411 | 0.022509 | 0.058566 | 0.027651 | 0.033912 | 0.027386 | 0.047081 | 0.038816 | 0.056214 | 0.043551 | 0.069474 | 0.1169 | 0.13528 | 0.036886 | 0.14081 | 0.1241 | 0.055498
| Rel.Std.Dev. |  | 1.818% | 1.175% | 2.187% | 2.293% | 2.649% | 5.006% | 0.8105% | 1.246% | 2.71% | 2.812% | 2.612% | 1.034% | 0.7292% | 1.955% | 0.8857% | 1.222% | 0.9526% | 1.576% | 1.365% | 1.919% | 1.404% | 2.196% | 3.854% | 4.188% | 1.238% | 4.177% | 4.029% | 1.734%
| PDM3 |  | -1.082% | 1.324% | 2.402% | 6.164% | 9.921% | -2.559% | 2.568% | -48.09% | -6.462% | 1.054% | 5.595% | -1.488% | 0.1345% | -2.839% | 1.278% | -9.976% | -6.732% | -3.056% | -7.759% | -4.948% | 0.6211% | 2.622% | -1.596% | 4.784% | -3.364% | 9.353% | -0.08178% | 3.811%
| Z-Score (Absolute) |  | 0.25 | 0.29 | 0.54 | 1.39 | 2.24 | 0.59 | 0.58 | 10.91 | 1.47 | 0.23 | 1.26 | 0.34 | 0.02 | 0.65 | 0.28 | 2.27 | 1.53 | 0.7 | 1.77 | 1.13 | 0.13 | 0.59 | 0.37 | 1.08 | 0.77 | 2.11 | 0.02 | 0.86
| Analytical results for Ga in OREAS 284 (Certified Value 9.76 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 10.1049 | 9.86 | 11.7 | 8.9744 | 8 | 7.8 | 9.46 | 10.45 | 9.94 | 11.05 | 9.45 | 9.1578 | 11 | 10.1 | 9.5 | 9.82 | 8.7 | 9.9 | 9.48 | 6.17 | 55.5 | 9 | [7 empty] | 1
| 1 | 2 | 9.9641 | 10.31 | 11 | 9.6314 | 8 | 8.4 | 9.65 | 10.2 | 9.74 | 11.1 | 9.82 | 9.399 | 11 | 10.6 | 9.9 | 9.81 | 8.9 | 9.5 | 10.39 | 5.9 | 52.5 | 9.5 | [7 empty] | 36
| 1 | 3 | 9.699 | 9.73 | 11 | 9.3514 | 8 | 9 | 9.73 | 9 | 9.65 | 10.85 | 9.3 | 9.316 | 11 | 10.3 | 9.6 | 9.82 | 8.4 | 9.7 | 10.61 | 5.98 | 50.6 | 9.4 | [7 empty] | 16
| 1 | 4 | 10.0016 | 9.65 | 11.4 | 9.0891 | 8 | 9 | 9.7 | 10.55 | 9.54 | 11.5 | 9.1 | 9.5321 | 11 | 10.5 | 9.8 | 9.67 | 9 | 9.7 | 10.2 | 5.95 | 53.6 | 9.4 | [7 empty] | 9.7618
| 1 | 5 | 9.9027 | 9.99 | 11.6 | 9.7378 | 8 | 9 | 9.77 | 10.4 | 10.15 | 11.25 | 9.49 | 9.0905 | 11 | 10.3 | 9.8 | 9.87 | 9 | 10 | 10.58 | 5.94 | 47.3 | 9.4 | [7 empty] | 90
| 1 | 6 | 10.2364 | 9.79 | 10.9 | 9.387 | 8 | 8.7 | 9.87 | 9.48 | 9.87 | 11.1 | 9.49 | 9.3189 | 11 | 10.7 | 9.7 | 9.74 | 8.9 | 9.8 | 10.5 | 5.78 | 51.2 | 9.5
| Mean |  | 9.9848 | 9.8883 | 11.2667 | 9.3619 | 8 | 8.65 | 9.6967 | 10.0133 | 9.815 | 11.1417 | 9.4417 | 9.3024 | 11 | 10.4167 | 9.7167 | 9.7883 | 8.8167 | 9.7667 | 10.2933 | 5.9533 | 51.7833 | 9.3667
| Median |  | 9.9828 | 9.825 | 11.2 | 9.3692 | 8 | 8.85 | 9.715 | 10.3 | 9.805 | 11.1 | 9.47 | 9.3174 | 11 | 10.4 | 9.75 | 9.815 | 8.9 | 9.75 | 10.445 | 5.945 | 51.85 | 9.4
| Std Dev. |  | 0.18266 | 0.23685 | 0.34448 | 0.29639 | 0 | 0.48062 | 0.13765 | 0.62838 | 0.21879 | 0.21775 | 0.23861 | 0.16021 | 0 | 0.22286 | 0.1472 | 0.071391 | 0.23166 | 0.17512 | 0.42538 | 0.12707 | 2.8096 | 0.18619
| Rel.Std.Dev. |  | 1.829% | 2.395% | 3.058% | 3.166% | 0% | 5.556% | 1.42% | 6.275% | 2.229% | 1.954% | 2.527% | 1.722% | 0% | 2.139% | 1.515% | 0.7293% | 2.628% | 1.793% | 4.133% | 2.134% | 5.426% | 1.988%
| PDM3 |  | 2.285% | 1.297% | 15.42% | -4.097% | -18.05% | -11.39% | -0.6669% | 2.577% | 0.5453% | 14.14% | -3.279% | -4.706% | 12.68% | 6.709% | -0.462% | 0.2722% | -9.682% | 0.0502% | 5.445% | -39.01% | 430.5% | -4.047%
| Z-Score (Absolute) |  | 0.34 | 0.18 | 2.43 | 0.68 | NA | 1.84 | 0.13 | 0.38 | 0.06 | 2.23 | 0.55 | 0.78 | NA | 1.04 | 0.1 | 0.02 | 1.57 | 0.02 | 0.84 | 6.24 | 68.54 | 0.67
Results from laboratories 1.11 and 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 284 (Certified Value 4.05 ppm) | [30 empty] | Certified
Gd | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | [7 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.29 | [7 empty] | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*MS | [7 empty] | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.5g | [7 empty] | 2
| 1 | 1 | 1.03: 4.477; 1.04: 3.8; 1.09: 4.4; 1.10: 4.8875; 1.11: 2.84 | [7 empty] | 1
| 1 | 2 | 1.03: 4.6026; 1.04: 3.756; 1.09: 4.2; 1.10: 5.1745; 1.11: 2.71 | [7 empty] | 37
| 1 | 3 | 1.03: 4.2756; 1.04: 3.693; 1.09: 4.3; 1.10: 4.8679; 1.11: 2.79 | [7 empty] | 16
| 1 | 4 | 1.03: 4.1896; 1.04: 3.613; 1.09: 4.1; 1.10: 4.9954; 1.11: 2.74 | [7 empty] | 4.0522
| 1 | 5 | 1.03: 4.4721; 1.04: 3.698; 1.09: 4.4; 1.10: 5.2693; 1.11: 2.7 | [7 empty] | 91
| 1 | 6 | 1.03: 4.5989; 1.04: 3.868; 1.09: 4.3; 1.10: 5.2981; 1.11: 2.55
| Mean |  | 1.03: 4.436; 1.04: 3.738; 1.09: 4.2833; 1.10: 5.0821; 1.11: 2.7217
| Median |  | 1.03: 4.4746; 1.04: 3.727; 1.09: 4.3; 1.10: 5.085; 1.11: 2.725
| Std Dev. |  | 1.03: 0.16953; 1.04: 0.089798; 1.09: 0.1169; 1.10: 0.19053; 1.11: 0.099079
| Rel.Std.Dev. |  | 1.03: 3.822%; 1.04: 2.402%; 1.09: 2.729%; 1.10: 3.749%; 1.11: 3.64%
| PDM3 |  | 1.03: 9.47%; 1.04: -7.754%; 1.09: 5.703%; 1.10: 25.42%; 1.11: -32.84%
| Z-Score (Absolute) |  | 1.03: 0.19; 1.04: 0.67; 1.09: 0; 1.10: 0.99; 1.11: 1.93
| Analytical results for Ge in OREAS 284 (Certified Value 0.12 ppm) | [30 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.29 | 1.31 | 1.32 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [15 empty] | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | [15 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.1 | 0.12 | 0.14 | 0.1 | 0.14 | 1.4 | 0.2 | 0.18 | 0.1 | 0.05 | 0.11 | < 0.1 | [15 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.1 | 0.11 | 0.15 | 0.09 | 0.17 | 1.4 | 0.2 | 0.17 | 0.1 | 0.03 | 0.1 | < 0.1 | [15 empty] | 2
| 1 | 3 | < 0.2 | < 0.1 | 0.12 | 0.1 | 0.13 | 0.1 | 0.16 | 1.4 | 0.2 | 0.19 | 0.1 | 0.05 | 0.1 | < 0.1 | [15 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.12 | 0.12 | 0.14 | 0.11 | 0.15 | 1.4 | 0.1 | 0.18 | 0.1 | 0.04 | 0.11 | < 0.1 | [15 empty] | 0.12194
| 1 | 5 | < 0.2 | < 0.1 | 0.12 | 0.13 | 0.14 | 0.11 | 0.19 | 1.4 | 0.2 | 0.18 | 0.1 | 0.05 | 0.1 | < 0.1 | [15 empty] | 92
| 1 | 6 | < 0.2 | < 0.1 | 0.1 | 0.11 | 0.14 | 0.1 | 0.16 | 1.4 | 0.1 | 0.19 | 0.1 | 0.04 | 0.1 | < 0.1
| Mean |  | 1.09: 0.11; 1.10: 0.115; 1.11: 0.14; 1.13: 0.10167; 1.14: 0.16167; 1.15: 1.4; 1.16: 0.16667; 1.17: 0.18167; 1.18: 0.1; 1.19: 0.043333; 1.20: 0.10333
| Median |  | 1.09: 0.11; 1.10: 0.115; 1.11: 0.14; 1.13: 0.1; 1.14: 0.16; 1.15: 1.4; 1.16: 0.2; 1.17: 0.18; 1.18: 0.1; 1.19: 0.045; 1.20: 0.1
| Std Dev. |  | 1.09: 0.010954; 1.10: 0.010488; 1.11: 0.0063246; 1.13: 0.0075277; 1.14: 0.017224; 1.15: 2.4324e-16; 1.16: 0.05164; 1.17: 0.0075277; 1.18: 1.5202e-17; 1.19: 0.008165; 1.20: 0.005164
| Rel.Std.Dev. |  | 1.09: 9.959%; 1.10: 9.12%; 1.11: 4.518%; 1.13: 7.404%; 1.14: 10.65%; 1.15: 1.737e-14%; 1.16: 30.98%; 1.17: 4.144%; 1.18: 1.52e-14%; 1.19: 18.84%; 1.20: 4.997%
| PDM3 |  | 1.09: -9.795%; 1.10: -5.695%; 1.11: 14.81%; 1.13: -16.63%; 1.14: 32.57%; 1.15: 1048%; 1.16: 36.67%; 1.17: 48.97%; 1.18: -18%; 1.19: -64.46%; 1.20: -15.26%
| Z-Score (Absolute) |  | 0.06 | 0.98 | 0.12 | 0.21 | 0.67 | 0.03 | 1.07 | NA | NA | 1.44 | NA | 1.1 | 0 | 0.98
Results from laboratories 1.20, 1.21 and 1.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 284 (Certified Value 0.34 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [11 empty] | 2
| 1 | 1 | 0.36508 | 0.28 | 0.3 | 0.28 | 0.32 | 0.34 | 0.34 | 0.28 | 0.31723 | 0.3 | 0.31 | 0.37 | 0.31 | 0.46 | 0.33 | 0.38 | 0.4 | 0.44 | [11 empty] | 1
| 1 | 2 | 0.35444 | 0.29 | 0.3 | 0.3 | 0.32 | 0.34 | 0.31 | 0.29 | 0.30045 | 0.3 | 0.31 | 0.37 | 0.3 | 0.42 | 0.37 | 0.38 | 0.37 | 0.45 | [11 empty] | 22
| 1 | 3 | 0.35196 | 0.28 | 0.3 | 0.33 | 0.27 | 0.35 | 0.33 | 0.27 | 0.31561 | 0.3 | 0.3 | 0.37 | 0.31 | 0.41 | 0.36 | 0.39 | 0.35 | 0.4 | [11 empty] | 16
| 1 | 4 | 0.35127 | 0.28 | 0.3 | 0.33 | 0.31 | 0.32 | 0.32 | 0.28 | 0.34493 | 0.3 | 0.31 | 0.37 | 0.3 | 0.4 | 0.34 | 0.41 | 0.38 | 0.44 | [11 empty] | 0.34175
| 1 | 5 | 0.3565 | 0.28 | 0.3 | 0.33 | 0.32 | 0.36 | 0.36 | 0.28 | 0.33673 | 0.3 | 0.3 | 0.37 | 0.3 | 0.4 | 0.38 | 0.38 | 0.32 | 0.41 | [11 empty] | 93
| 1 | 6 | 0.35889 | 0.25 | 0.3 | 0.3 | 0.35 | 0.34 | 0.3 | 0.29 | 0.32496 | 0.3 | 0.3 | 0.38 | 0.31 | 0.39 | 0.36 | 0.41 | 0.33 | 0.43
| Mean |  | 0.35636 | 0.27667 | 0.3 | 0.31167 | 0.315 | 0.34167 | 0.32667 | 0.28167 | 0.32332 | 0.3 | 0.305 | 0.37167 | 0.305 | 0.41333 | 0.35667 | 0.39167 | 0.35833 | 0.42833
| Median |  | 0.35547 | 0.28 | 0.3 | 0.315 | 0.32 | 0.34 | 0.325 | 0.28 | 0.3211 | 0.3 | 0.305 | 0.37 | 0.305 | 0.405 | 0.36 | 0.385 | 0.36 | 0.435
| Std Dev. |  | 0.0051263 | 0.013663 | 0 | 0.02137 | 0.025884 | 0.013292 | 0.021602 | 0.0075277 | 0.01593 | 0 | 0.0054772 | 0.0040825 | 0.0054772 | 0.025033 | 0.018619 | 0.01472 | 0.030605 | 0.019408
| Rel.Std.Dev. |  | 1.439% | 4.938% | 0% | 6.857% | 8.217% | 3.89% | 6.613% | 2.673% | 4.927% | 0% | 1.796% | 1.098% | 1.796% | 6.056% | 5.22% | 3.758% | 8.541% | 4.531%
| PDM3 |  | 4.274% | -19.04% | -12.22% | -8.803% | -7.828% | -0.02455% | -4.414% | -17.58% | -5.393% | -12.22% | -10.75% | 8.754% | -10.75% | 20.95% | 4.365% | 14.61% | 4.852% | 25.34%
| Z-Score (Absolute) |  | 0.56 | 1.45 | NA | 0.57 | 0.48 | 0.19 | 0.19 | 1.33 | 0.27 | NA | 0.74 | 0.95 | 0.74 | 2 | 0.57 | 1.45 | 0.61 | 2.38
Results from laboratories 1.09 and 1.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 284 (Indicative Value 0.014 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.29 | 1.30 | 1.31 | 1.32 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [8 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 15g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | [8 empty] | 3
| 1 | 1 | < 1 | < 0.1 | < 0.01 | < 5 | 0.08 | 0.066 | < 0.01 | < 0.01 | < 0.01 | 0.02 | 0.01 | < 0.05 | 0.851 | 0.02 | 0.12 | < 0.005 | < 0.01 | 0.22 | < 1 | < 0.01 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 1 | < 0.1 | < 0.01 | < 5 | 0.07 | 0.086 | 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.05 | 0.824 | 0.02 | 0.11 | < 0.005 | < 0.01 | 0.21 | < 1 | < 0.01 | < 0.01 | [8 empty] | 5
| 1 | 3 | < 1 | < 0.1 | < 0.01 | < 5 | 0.07 | 0.062 | 0.01 | 0.01 | < 0.01 | 0.03 | 0.01 | < 0.05 | 0.818 | 0.01 | 0.12 | < 0.005 | < 0.01 | 0.22 | < 1 | < 0.01 | < 0.01 | [8 empty] | 16
| 1 | 4 | < 1 | < 0.1 | < 0.01 | < 5 | 0.07 | 0.057 | < 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | < 0.05 | 0.859 | < 0.01 | 0.12 | < 0.005 | < 0.01 | 0.22 | < 1 | < 0.01 | < 0.01 | [8 empty] | 0.013738
| 1 | 5 | < 1 | < 0.1 | < 0.01 | < 5 | 0.07 | 0.04 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | 0.818 | < 0.01 | 0.13 | < 0.005 | < 0.01 | 0.22 | < 1 | < 0.01 | < 0.01 | [8 empty] | 11
| 1 | 6 | < 1 | < 0.1 | 0.02 | < 5 | 0.07 | 0.057 | 0.02 | 0.01 | 0.01 | 0.02 | < 0.01 | < 0.05 | 0.785 | 0.01 | 0.12 | < 0.005 | < 0.01 | 0.21 | < 1 | < 0.01 | < 0.01
| Mean |  | 1.09: 0.02; 1.12: 0.071667; 1.13: 0.061333; 1.14: 0.0125; 1.15: 0.012; 1.16: 0.01; 1.17: 0.016667; 1.18: 0.01; 1.20: 0.82583; 1.21: 0.015; 1.23: 0.12; 1.29: 0.21667
| Median |  | 1.09: 0.02; 1.12: 0.07; 1.13: 0.0595; 1.14: 0.01; 1.15: 0.01; 1.16: 0.01; 1.17: 0.015; 1.18: 0.01; 1.20: 0.821; 1.21: 0.015; 1.23: 0.12; 1.29: 0.22
| Std Dev. |  | 1.12: 0.0040825; 1.13: 0.014989; 1.14: 0.005; 1.15: 0.0044721; 1.16: 0; 1.17: 0.008165; 1.18: 0; 1.20: 0.026559; 1.21: 0.0057735; 1.23: 0.0063246; 1.29: 0.005164
| Rel.Std.Dev. |  | 1.12: 5.696%; 1.13: 24.44%; 1.14: 40%; 1.15: 37.27%; 1.16: 0%; 1.17: 48.99%; 1.18: 0%; 1.20: 3.216%; 1.21: 38.49%; 1.23: 5.27%; 1.29: 2.383%
| PDM3 |  | 1.09: 45.58%; 1.12: 421.7%; 1.13: 346.4%; 1.14: -9.012%; 1.15: -12.65%; 1.16: -27.21%; 1.17: 21.32%; 1.18: -27.21%; 1.20: 5911%; 1.21: 9.185%; 1.23: 773.5%; 1.29: 1477%
| Z-Score (Absolute) |  | 27.94 | 1.93 | 0.53 | 143.53 | 3.18 | 2.58 | 0.39 | 0.34 | 0.58 | 0 | 0.48 | 0.48 | 46.77 | 0.29 | 5.97 | 0.82 | 0.67 | 11.56 | 27.94 | 0.67 | 0.67
| Analytical results for Ho in OREAS 284 (Certified Value 0.41 ppm) | [30 empty] | Certified
Ho | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.11: Lab; 1.12: Lab | [8 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.11: 1.10; 1.12: 1.29 | [8 empty] | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.09: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS | [8 empty] | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.09: 0.5g; 1.11: 0.5g; 1.12: 0.5g | [8 empty] | 3
| 1 | 1 | 1.03: 0.44954; 1.04: 0.412; 1.09: 0.4; 1.11: 0.403; 1.12: 0.37 | [8 empty] | 1
| 1 | 2 | 1.03: 0.44142; 1.04: 0.399; 1.09: 0.4; 1.11: 0.4302; 1.12: 0.35 | [8 empty] | 23
| 1 | 3 | 1.03: 0.44814; 1.04: 0.409; 1.09: 0.4; 1.11: 0.4115; 1.12: 0.35 | [8 empty] | 16
| 1 | 4 | 1.03: 0.43849; 1.04: 0.404; 1.09: 0.4; 1.11: 0.4141; 1.12: 0.36 | [8 empty] | 0.40563
| 1 | 5 | 1.03: 0.44398; 1.04: 0.404; 1.09: 0.4; 1.11: 0.4378; 1.12: 0.35 | [8 empty] | 94
| 1 | 6 | 1.03: 0.44604; 1.04: 0.412; 1.09: 0.4; 1.11: 0.4446; 1.12: 0.34
| Mean |  | 1.03: 0.4446; 1.04: 0.40667; 1.09: 0.4; 1.11: 0.42353; 1.12: 0.35333
| Median |  | 1.03: 0.44501; 1.04: 0.4065; 1.09: 0.4; 1.11: 0.42215; 1.12: 0.35
| Std Dev. |  | 1.03: 0.0041685; 1.04: 0.0052026; 1.09: 6.0809e-17; 1.11: 0.016415; 1.12: 0.010328
| Rel.Std.Dev. |  | 1.03: 0.9376%; 1.04: 1.279%; 1.09: 1.52e-14%; 1.11: 3.876%; 1.12: 2.923%
| PDM3 |  | 1.03: 9.608%; 1.04: 0.2564%; 1.09: -1.387%; 1.11: 4.415%; 1.12: -12.89%
| Z-Score (Absolute) |  | 1.03: 1.52; 1.04: 0; 1.09: 0.27; 1.11: 0.67; 1.12: 2.13
| Analytical results for In in OREAS 284 (Certified Value 0.066 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.04 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.31 | 1.32 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | [12 empty] | 3
| 1 | 1 | 0.07 | 0.06 | < 5 | 0.066 | 0.073 | 0.067 | 0.058 | 0.066 | 0.063264 | 0.04 | 0.05 | 0.07 | 0.063 | 0.07 | 0.06 | 0.088 | 0.06 | [12 empty] | 1
| 1 | 2 | 0.07 | 0.07 | < 5 | 0.065 | 0.071 | 0.06 | 0.059 | 0.071 | 0.067939 | 0.04 | 0.06 | 0.07 | 0.062 | 0.06 | 0.07 | 0.078 | 0.06 | [12 empty] | 24
| 1 | 3 | 0.07 | 0.07 | < 5 | 0.067 | 0.064 | 0.069 | 0.057 | 0.064 | 0.062281 | 0.04 | 0.06 | 0.07 | 0.06 | 0.07 | 0.06 | 0.072 | 0.08 | [12 empty] | 16
| 1 | 4 | 0.06 | 0.07 | < 5 | 0.069 | 0.074 | 0.069 | 0.068 | 0.061 | 0.069734 | 0.04 | 0.06 | 0.07 | 0.068 | 0.07 | 0.06 | 0.084 | 0.06 | [12 empty] | 0.066433
| 1 | 5 | 0.07 | 0.07 | < 5 | 0.069 | 0.068 | 0.064 | 0.068 | 0.068 | 0.06699 | 0.04 | 0.05 | 0.06 | 0.062 | 0.06 | 0.07 | 0.067 | 0.07 | [12 empty] | 95
| 1 | 6 | 0.06 | 0.06 | < 5 | 0.062 | 0.066 | 0.066 | 0.062 | 0.059 | 0.066381 | 0.04 | 0.06 | 0.07 | 0.061 | 0.06 | 0.07 | 0.08 | 0.07
| Mean |  | 0.066667 | 0.066667 |  | 0.066333 | 0.069333 | 0.065833 | 0.062 | 0.064833 | 0.066098 | 0.04 | 0.056667 | 0.068333 | 0.062667 | 0.065 | 0.065 | 0.078167 | 0.066667
| Median |  | 0.07 | 0.07 |  | 0.0665 | 0.0695 | 0.0665 | 0.0605 | 0.065 | 0.066686 | 0.04 | 0.06 | 0.07 | 0.062 | 0.065 | 0.065 | 0.079 | 0.065
| Std Dev. |  | 0.005164 | 0.005164 |  | 0.0026583 | 0.0039833 | 0.0034303 | 0.0049396 | 0.004446 | 0.0028317 | 0 | 0.005164 | 0.0040825 | 0.0028048 | 0.0054772 | 0.0054772 | 0.007705 | 0.008165
| Rel.Std.Dev. |  | 7.746% | 7.746% |  | 4.008% | 5.745% | 5.211% | 7.967% | 6.858% | 4.284% | 0% | 9.113% | 5.974% | 4.476% | 8.427% | 8.427% | 9.857% | 12.25%
| PDM3 |  | 0.3515% | 0.3515% |  | -0.1503% | 4.366% | -0.9029% | -6.673% | -2.408% | -0.5042% | -39.79% | -14.7% | 2.86% | -5.67% | -2.157% | -2.157% | 17.66% | 0.3515%
| Z-Score (Absolute) |  | 0.3 | 0.3 | 1297.55 | 0.13 | 1.72 | 0.14 | 2.18 | 0.67 | 0 | 13.91 | 5.03 | 1.19 | 1.83 | 0.59 | 0.59 | 6.43 | 0.3
| Analytical results for K in OREAS 284 (Certified Value 1.02 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.99752 | 1.0399 | 1.0914 | 1.0044 | 0.99 | 0.9839 | 0.91 | 0.99 | 1.05 | 1.17 | 1.01 | 1.04 | 1.03 | 0.9576 | 0.93 | 1.04 | 1.05 | 0.98 | 1.03 | 0.97 | 0.9444 | 0.9458 | 1.27 | 1.1622 | 0.79216 | 1.31 | 1.04 |  |  | 1
| 1 | 2 | 0.9596 | 1.0268 | 1.0992 | 0.97647 | 0.96 | 1.0506 | 0.9 | 1.02 | 1.07 | 1.17 | 1.02 | 1.02 | 1.01 | 0.95463 | 0.94 | 1.05 | 1.03 | 0.99 | 1.04 | 0.94 | 0.9243 | 1.0032 | 1.272 | 1.1207 | 0.8072 | 1.22 | 1.07 |  |  | 0
| 1 | 3 | 0.96023 | 1.0332 | 1.0622 | 1.0045 | 0.96 | 0.9249 | 0.91 | 1.02 | 1.07 | 1.16 | 1 | 1.04 | 1.01 | 0.9409 | 0.94 | 1.03 | 1.06 | 0.98 | 1.05 | 0.96 | 0.9454 | 1.0382 | 1.356 | 1.1207 | 0.77784 | 1.3 | 1.07 |  |  | 16
| 1 | 4 | 0.98238 | 1.0504 | 1.0459 | 0.98726 | 0.99 | 0.9077 | 0.91 | 1.04 | 1.06 | 1.15 | 0.99 | 1.02 | 1.01 | 0.97416 | 0.95 | 1.04 | 1.05 | 0.97 | 1.05 | 0.98 | 0.9117 | 1.0042 | 1.25 | 1.1456 | 0.73584 | 1.25 | 1.07 |  |  | 1.0152
| 1 | 5 | 0.9789 | 1.0192 | 1.1038 | 0.99997 | 0.95 | 0.9355 | 0.9 | 1.03 | 1.08 | 1.16 | 1.02 | 1.04 | 1.03 | 0.98073 | 0.96 | 1.05 | 1.05 | 0.98 | 1.06 | 0.97 | 0.9308 | 1.0403 | 1.398 | 1.1456 | 0.8216 | 1.19 | 1.08 |  |  | 96
| 1 | 6 | 0.9697 | 1.0417 | 1.0552 | 1.005 | 0.87 | 0.9595 | 0.9 | 0.99 | 1 | 1.07 | 1.01 | 1.04 | 1.07 | 0.99025 | 0.95 | 1.03 | 1.05 | 0.96 | 1.05 | 0.97 | 0.8962 | 1.0132 | 1.311 | 1.129 | 0.8152 | 1.26 | 1.1
| Mean |  | 0.97472 | 1.0352 | 1.0763 | 0.99626 | 0.95333 | 0.96035 | 0.905 | 1.015 | 1.055 | 1.1467 | 1.0083 | 1.0333 | 1.0267 | 0.96638 | 0.945 | 1.04 | 1.0483 | 0.97667 | 1.0467 | 0.965 | 0.92547 | 1.0075 | 1.3095 | 1.1373 | 0.79164 | 1.255 | 1.0717
| Median |  | 0.9743 | 1.0366 | 1.0768 | 1.0022 | 0.96 | 0.9475 | 0.905 | 1.02 | 1.065 | 1.16 | 1.01 | 1.04 | 1.02 | 0.96588 | 0.945 | 1.04 | 1.05 | 0.98 | 1.05 | 0.97 | 0.92755 | 1.0087 | 1.2915 | 1.1373 | 0.79968 | 1.255 | 1.07
| Std Dev. |  | 0.014559 | 0.011189 | 0.024821 | 0.011805 | 0.044121 | 0.051666 | 0.0054772 | 0.020736 | 0.02881 | 0.038297 | 0.01169 | 0.010328 | 0.023381 | 0.018442 | 0.010488 | 0.0089443 | 0.0098319 | 0.010328 | 0.010328 | 0.013784 | 0.01914 | 0.034314 | 0.057452 | 0.0166 | 0.03162 | 0.045935 | 0.019408
| Rel.Std.Dev. |  | 1.494% | 1.081% | 2.306% | 1.185% | 4.628% | 5.38% | 0.6052% | 2.043% | 2.731% | 3.34% | 1.159% | 0.9995% | 2.277% | 1.908% | 1.11% | 0.86% | 0.9379% | 1.057% | 0.9867% | 1.428% | 2.068% | 3.406% | 4.387% | 1.46% | 3.994% | 3.66% | 1.811%
| PDM3 |  | -3.985% | 1.973% | 6.019% | -1.862% | -6.091% | -5.4% | -10.85% | -0.01686% | 3.923% | 12.95% | -0.6736% | 1.789% | 1.132% | -4.806% | -6.912% | 2.446% | 3.267% | -3.793% | 3.102% | -4.942% | -8.836% | -0.7573% | 28.99% | 12.03% | -22.02% | 23.62% | 5.565%
| Z-Score (Absolute) |  | 0.56 | 0.28 | 0.85 | 0.26 | 0.86 | 0.76 | 1.53 | 0 | 0.55 | 1.83 | 0.09 | 0.25 | 0.16 | 0.67 | 0.97 | 0.35 | 0.46 | 0.53 | 0.44 | 0.69 | 1.24 | 0.1 | 4.08 | 1.7 | 3.1 | 3.33 | 0.79
| Analytical results for La in OREAS 284 (Certified Value 19.2 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 1
| 1 | 1 | 19.2103 | 18.305 | 22 | 23.2128 | 19 | 17.8 | 19.2 | 22.1 | 17.9 | 17.1 | 17 | 23.0534 | 15 | 15.5 | 26 | 15.1 | 13.2 | 19.4 | 30 | 21.098 | 8 | 14.56 | 27 | 19.2 | [5 empty] | 1
| 1 | 2 | 19.0926 | 17.477 | 20.7 | 23.7297 | 19 | 19.2 | 19.2 | 22.2 | 17.8 | 16.4 | 17.6 | 22.5844 | 16 | 15.9 | 26.4 | 15 | 13.9 | 19.6 | 30 | 23.336 | 8 | 14.7264 | 25.7 | 19.8 | [5 empty] | 25
| 1 | 3 | 19.0267 | 17.263 | 21.1 | 23.0758 | 19 | 19.3 | 19.7 | 19.7 | 17.2 | 16.8 | 16.4 | 19.3812 | 16 | 15.8 | 26.3 | 15.3 | 12.3 | 19.8 | 31 | 24.104 | 8 | 14.0736 | 24.3 | 19.9 | [5 empty] | 16
| 1 | 4 | 18.3933 | 16.652 | 20.4 | 22.2258 | 18 | 19.4 | 20 | 23.2 | 16.4 | 17.9 | 16.7 | 23.5256 | 16 | 15.8 | 27.6 | 14.7 | 13 | 19.9 | 30 | 22.66 | 8 | 14.7456 | 26.2 | 20.8 | [5 empty] | 19.2174
| 1 | 5 | 19.7523 | 16.94 | 22 | 24.2918 | 19 | 19.2 | 20.1 | 21.9 | 17.8 | 17.4 | 17.6 | 20.0256 | 17 | 15.3 | 27.3 | 15.1 | 14.1 | 21.4 | 31 | 23.232 | 8 | 14.5536 | 22.1 | 20.1 | [5 empty] | 97
| 1 | 6 | 19.8209 | 14.787 | 21.2 | 23.1638 | 18 | 17 | 16.5 | 19 | 17.7 | 16.8 | 17.3 | 21.92 | 17 | 15.6 | 27.3 | 15.3 | 13.1 | 21.7 | 30 | 24.008 | 8 | 14.8864 | 24.1 | 20.6
| Mean |  | 19.216 | 16.904 | 21.2333 | 23.2833 | 18.6667 | 18.65 | 19.1167 | 21.35 | 17.4667 | 17.0667 | 17.1 | 21.7484 | 16.1667 | 15.65 | 26.8167 | 15.0833 | 13.2667 | 20.3 | 30.3333 | 23.073 | 8 | 14.5909 | 24.9 | 20.0667
| Median |  | 19.1515 | 17.1015 | 21.15 | 23.1883 | 19 | 19.2 | 19.45 | 22 | 17.75 | 16.95 | 17.15 | 22.2522 | 16 | 15.7 | 26.85 | 15.1 | 13.15 | 19.85 | 30 | 23.284 | 8 | 14.6432 | 25 | 20
| Std Dev. |  | 0.52562 | 1.1804 | 0.65929 | 0.69238 | 0.5164 | 1.0035 | 1.3378 | 1.6282 | 0.5785 | 0.52789 | 0.4899 | 1.6829 | 0.75277 | 0.22583 | 0.66156 | 0.22286 | 0.6532 | 0.98793 | 0.5164 | 1.1048 | 0 | 0.28269 | 1.7652 | 0.5785
| Rel.Std.Dev. |  | 2.735% | 6.983% | 3.105% | 2.974% | 2.766% | 5.381% | 6.998% | 7.626% | 3.312% | 3.093% | 2.865% | 7.738% | 4.656% | 1.443% | 2.467% | 1.478% | 4.924% | 4.867% | 1.702% | 4.788% | 0% | 1.937% | 7.089% | 2.883%
| PDM3 |  | -0.007049% | -12.04% | 10.49% | 21.16% | -2.866% | -2.952% | -0.5241% | 11.1% | -9.11% | -11.19% | -11.02% | 13.17% | -15.87% | -18.56% | 39.54% | -21.51% | -30.97% | 5.634% | 57.84% | 20.06% | -58.37% | -24.07% | 29.57% | 4.419%
| Z-Score (Absolute) |  | 0.08 | 0.52 | 0.61 | 1.14 | 0.06 | 0.06 | 0.06 | 0.64 | 0.37 | 0.47 | 0.47 | 0.74 | 0.71 | 0.84 | 2.06 | 0.99 | 1.46 | 0.37 | 2.98 | 1.09 | 2.83 | 1.12 | 1.56 | 0.31
| Analytical results for Li in OREAS 284 (Certified Value 44.4 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 45.3978 | 43.8 | 55.6529 | 46.408 | 49.8 | 54.6188 | 46 | 45.9 | 48.1 | 44 | 45.6 | 45.2 | 39.3113 | 40.5 | 46 | 42.3 | 39.9 | 44 | 45 | 42.3 | 49.62 | 39 | 103.84 | 46.4 | 42 |  |  |  |  | 1
| 1 | 2 | 43.4644 | 44 | 55.29 | 45.9325 | 46.2 | 57.0039 | 46 | 46.5 | 48.5 | 44.4 | 45.3 | 44.8 | 42.3208 | 40.8 | 46 | 42 | 39.8 | 42 | 44 | 42.8 | 51.61 | 38 | 102.8 | 45.5 | 44 |  |  |  |  | 45
| 1 | 3 | 43.586 | 43.8 | 53.9128 | 46.6134 | 46.1 | 47.9951 | 46 | 46.3 | 47.3 | 43.7 | 45.4 | 44.8 | 41.3198 | 41 | 44 | 42.7 | 40.5 | 43 | 45 | 44.9 | 52.72 | 38 | 99.92 | 44.2 | 44 |  |  |  |  | 16
| 1 | 4 | 44.2725 | 43.2 | 52.6914 | 44.4204 | 46.1 | 50.9075 | 45 | 46.2 | 46.9 | 43.2 | 46.3 | 44.8 | 44.4345 | 39.6 | 45 | 41.9 | 39 | 45 | 43 | 42.9 | 51.48 | 39 | 94.72 | 46.9 | 43 |  |  |  |  | 44.4363
| 1 | 5 | 43.7233 | 43.6 | 56.6241 | 46.4563 | 48.6 | 49.1258 | 45 | 46.8 | 47.7 | 44.3 | 46.4 | 45.4 | 42.5262 | 40 | 45 | 42.7 | 39.5 | 44 | 44 | 46.2 | 53.41 | 38 | 102.32 | 43.1 | 44 |  |  |  |  | 98
| 1 | 6 | 44.5775 | 43.8 | 53.877 | 45.6129 | 45.2 | 48.3966 | 45 | 42.4 | 45.7 | 44.4 | 45.7 | 46.7 | 43.8993 | 43.2 | 45 | 42.1 | 38.7 | 44 | 43 | 43.2 | 49.95 | 38 | 99.92 | 43.3 | 45
| Mean |  | 44.1703 | 43.7 | 54.6747 | 45.9072 | 47 | 51.3413 | 45.5 | 45.6833 | 47.3667 | 44 | 45.7833 | 45.2833 | 42.302 | 40.85 | 45.1667 | 42.2833 | 39.5667 | 43.6667 | 44 | 43.7167 | 51.465 | 38.3333 | 100.5867 | 44.9 | 43.6667
| Median |  | 43.9979 | 43.8 | 54.6014 | 46.1703 | 46.15 | 50.0166 | 45.5 | 46.25 | 47.5 | 44.15 | 45.65 | 45 | 42.4235 | 40.65 | 45 | 42.2 | 39.65 | 44 | 44 | 43.05 | 51.545 | 38 | 101.12 | 44.85 | 44
| Std Dev. |  | 0.7379 | 0.27568 | 1.4343 | 0.81838 | 1.7833 | 3.6816 | 0.54772 | 1.6364 | 0.99331 | 0.47749 | 0.46224 | 0.73869 | 1.8471 | 1.2613 | 0.75277 | 0.34881 | 0.65013 | 1.0328 | 0.89443 | 1.5065 | 1.4885 | 0.5164 | 3.2801 | 1.6062 | 1.0328
| Rel.Std.Dev. |  | 1.671% | 0.6308% | 2.623% | 1.783% | 3.794% | 7.171% | 1.204% | 3.582% | 2.097% | 1.085% | 1.01% | 1.631% | 4.366% | 3.088% | 1.667% | 0.8249% | 1.643% | 2.365% | 2.033% | 3.446% | 2.892% | 1.347% | 3.261% | 3.577% | 2.365%
| PDM3 |  | -0.5987% | -1.657% | 23.04% | 3.31% | 5.769% | 15.54% | 2.394% | 2.806% | 6.595% | -0.9818% | 3.031% | 1.906% | -4.803% | -8.071% | 1.644% | -4.845% | -10.96% | -1.732% | -0.9818% | -1.619% | 15.82% | -13.73% | 126.4% | 1.044% | -1.732%
| Z-Score (Absolute) |  | 0.4 | 0.66 | 5.34 | 0.55 | 1.15 | 3.52 | 0.33 | 0.43 | 1.35 | 0.49 | 0.48 | 0.21 | 1.42 | 2.21 | 0.15 | 1.43 | 2.92 | 0.67 | 0.49 | 0.65 | 3.59 | 3.59 | 30.45 | 0 | 0.67
| Analytical results for Lu in OREAS 284 (Certified Value 0.095 ppm) | [30 empty] | Certified
Lu | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.21; 1.10: 1.25; 1.11: 1.32 |  |  |  |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*OES/MS |  |  |  |  | 3
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.25g; 1.10: 0.25g; 1.11: 0.25g |  |  |  |  | 3
| 1 | 1 | 1.03: 0.10157; 1.04: 0.09; 1.05: < 0.1; 1.07: 0.0889; 1.09: 0.1; 1.10: 0.1; 1.11: 0.08 |  |  |  |  | 1
| 1 | 2 | 1.03: 0.10068; 1.04: 0.09; 1.05: < 0.1; 1.07: 0.0918; 1.09: 0.11; 1.10: 0.09; 1.11: 0.09 |  |  |  |  | 26
| 1 | 3 | 1.03: 0.098899; 1.04: 0.094; 1.05: < 0.1; 1.07: 0.0841; 1.09: 0.11; 1.10: 0.1; 1.11: 0.09 |  |  |  |  | 16
| 1 | 4 | 1.03: 0.099155; 1.04: 0.087; 1.05: < 0.1; 1.07: 0.0902; 1.09: 0.1; 1.10: 0.1; 1.11: 0.09 |  |  |  |  | 0.095156
| 1 | 5 | 1.03: 0.10013; 1.04: 0.09; 1.05: < 0.1; 1.07: 0.0928; 1.09: 0.1; 1.10: 0.1; 1.11: 0.09 |  |  |  |  | 99
| 1 | 6 | 1.03: 0.101; 1.04: 0.091; 1.05: < 0.1; 1.07: 0.0944; 1.09: 0.1; 1.10: 0.1; 1.11: 0.09
| Mean |  | 1.03: 0.10024; 1.04: 0.090333; 1.07: 0.090367; 1.09: 0.10333; 1.10: 0.098333; 1.11: 0.088333
| Median |  | 1.03: 0.1004; 1.04: 0.09; 1.07: 0.091; 1.09: 0.1; 1.10: 0.1; 1.11: 0.09
| Std Dev. |  | 1.03: 0.0010513; 1.04: 0.0022509; 1.07: 0.0036247; 1.09: 0.005164; 1.10: 0.0040825; 1.11: 0.0040825
| Rel.Std.Dev. |  | 1.03: 1.049%; 1.04: 2.492%; 1.07: 4.011%; 1.09: 4.997%; 1.10: 4.152%; 1.11: 4.622%
| PDM3 |  | 1.03: 5.341%; 1.04: -5.069%; 1.07: -5.034%; 1.09: 8.593%; 1.10: 3.339%; 1.11: -7.17%
| Z-Score (Absolute) |  | 1.03: 0.84; 1.04: 0; 1.05: 3.42; 1.07: 0; 1.09: 1.1; 1.10: 0.67; 1.11: 0.17
| Analytical results for Mg in OREAS 284 (Certified Value 0.81 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.829 | 0.83 | 0.83711 | 0.85523 | 0.83 | 0.8069 | 0.76 | 0.76 | 0.82 | 0.9 | 0.79 | 0.81 | 0.81 | 0.80667 | 0.78 | 0.76 | 0.81 | 0.79 | 0.8 | 0.81 | 0.82 | 0.74 | 0.928 | 0.862 | 0.8911 | 0.848 | 0.84 |  |  | 1
| 1 | 2 | 0.79894 | 0.82 | 0.84894 | 0.80363 | 0.78 | 0.8556 | 0.75 | 0.81 | 0.83 | 0.9 | 0.8 | 0.8 | 0.81 | 0.79376 | 0.78 | 0.77 | 0.85 | 0.8 | 0.81 | 0.78 | 0.81 | 0.78 | 0.913 | 0.838 | 0.9274 | 0.76 | 0.86 |  |  | 0
| 1 | 3 | 0.80035 | 0.82 | 0.82472 | 0.81987 | 0.78 | 0.7589 | 0.74 | 0.81 | 0.82 | 0.87 | 0.79 | 0.8 | 0.8 | 0.78641 | 0.79 | 0.75 | 0.87 | 0.78 | 0.74 | 0.8 | 0.83 | 0.81 | 0.982 | 0.832 | 0.8484 | 0.81 | 0.87 |  |  | 16
| 1 | 4 | 0.82166 | 0.84 | 0.80696 | 0.84297 | 0.8 | 0.7515 | 0.76 | 0.81 | 0.82 | 0.87 | 0.78 | 0.81 | 0.81 | 0.81374 | 0.8 | 0.76 | 0.88 | 0.78 | 0.81 | 0.81 | 0.79 | 0.78 | 0.915 | 0.85 | 0.8238 | 0.81 | 0.87 |  |  | 0.80997
| 1 | 5 | 0.81238 | 0.81 | 0.85148 | 0.82952 | 0.8 | 0.7696 | 0.76 | 0.81 | 0.83 | 0.89 | 0.8 | 0.82 | 0.82 | 0.81665 | 0.8 | 0.76 | 0.86 | 0.79 | 0.81 | 0.8 | 0.81 | 0.82 |  | 0.838 | 0.8544 | 0.8 | 0.87 |  |  | 100
| 1 | 6 | 0.82147 | 0.83 | 0.81138 | 0.82362 | 0.75 | 0.7853 | 0.75 | 0.77 | 0.77 | 0.84 | 0.79 | 0.8 | 0.84 | 0.82858 | 0.8 | 0.76 | 0.86 | 0.78 | 0.81 | 0.81 | 0.78 | 0.78 | 0.962 | 0.85 | 0.8694 | 0.801 | 0.89
| Mean |  | 0.81397 | 0.825 | 0.8301 | 0.82914 | 0.79 | 0.78797 | 0.75333 | 0.795 | 0.815 | 0.87833 | 0.79167 | 0.80667 | 0.815 | 0.80763 | 0.79167 | 0.76 | 0.855 | 0.78667 | 0.79667 | 0.80167 | 0.80667 | 0.785 | 0.94 | 0.845 | 0.86908 | 0.80483 | 0.86667
| Median |  | 0.81693 | 0.825 | 0.83091 | 0.82657 | 0.79 | 0.77745 | 0.755 | 0.81 | 0.82 | 0.88 | 0.79 | 0.805 | 0.81 | 0.8102 | 0.795 | 0.76 | 0.86 | 0.785 | 0.81 | 0.805 | 0.81 | 0.78 | 0.928 | 0.844 | 0.8619 | 0.8055 | 0.87
| Std Dev. |  | 0.012288 | 0.010488 | 0.018847 | 0.018106 | 0.026833 | 0.038605 | 0.008165 | 0.023452 | 0.022583 | 0.023166 | 0.0075277 | 0.008165 | 0.013784 | 0.015502 | 0.0098319 | 0.0063246 | 0.02429 | 0.008165 | 0.028048 | 0.01169 | 0.018619 | 0.028107 | 0.030602 | 0.011009 | 0.036276 | 0.02816 | 0.01633
| Rel.Std.Dev. |  | 1.51% | 1.271% | 2.27% | 2.184% | 3.397% | 4.899% | 1.084% | 2.95% | 2.771% | 2.638% | 0.9509% | 1.012% | 1.691% | 1.919% | 1.242% | 0.8322% | 2.841% | 1.038% | 3.521% | 1.458% | 2.308% | 3.581% | 3.256% | 1.303% | 4.174% | 3.499% | 1.884%
| PDM3 |  | 0.4928% | 1.855% | 2.485% | 2.366% | -2.466% | -2.717% | -6.993% | -1.849% | 0.6205% | 8.44% | -2.26% | -0.4083% | 0.6205% | -0.2889% | -2.26% | -6.17% | 5.559% | -2.878% | -1.643% | -1.026% | -0.4083% | -3.083% | 16.05% | 4.324% | 7.298% | -0.6347% | 6.999%
| Z-Score (Absolute) |  | 0.27 | 0.67 | 0.86 | 0.83 | 0.61 | 0.69 | 1.96 | 0.43 | 0.31 | 2.64 | 0.55 | 0 | 0.31 | 0.04 | 0.55 | 1.72 | 1.78 | 0.74 | 0.37 | 0.18 | 0 | 0.8 | 4.9 | 1.42 | 2.3 | 0.07 | 2.21
| Analytical results for Mn in OREAS 284 (Certified Value 0.033 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.033326 | 0.0319 | 0.033613 | 0.033105 | 0.0385 | 0.0327 | 0.0382 | 0.0309 | 0.0324 | 0.0348 | 0.0321 | 0.0339 | 0.0325 | 0.031994 | 0.0303 | 0.0311 | 0.0305 | 0.0326 | 0.0317 | 0.0329 | 0.0333 | 0.0313 | 0.03 | 0.0341 | 0.04161 | 0.0398 | 0.034 |  |  | 1
| 1 | 2 | 0.032465 | 0.0316 | 0.033254 | 0.032503 | 0.0372 | 0.0345 | 0.0385 | 0.0325 | 0.0327 | 0.0346 | 0.0326 | 0.0336 | 0.0325 | 0.034216 | 0.0311 | 0.0328 | 0.0317 | 0.0331 | 0.0323 | 0.0318 | 0.0326 | 0.0331 | 0.031 | 0.0325 | 0.04068 | 0.0378 | 0.0348 |  |  | 27
| 1 | 3 | 0.032028 | 0.0317 | 0.03211 | 0.033547 | 0.0372 | 0.0306 | 0.0381 | 0.0326 | 0.0328 | 0.0336 | 0.0319 | 0.0339 | 0.0318 | 0.03241 | 0.0302 | 0.0324 | 0.0328 | 0.0324 | 0.0318 | 0.0326 | 0.0334 | 0.0341 | 0.031 | 0.0325 | 0.03974 | 0.0369 | 0.035 |  |  | 16
| 1 | 4 | 0.033229 | 0.0323 | 0.031193 | 0.032899 | 0.0359 | 0.0302 | 0.0379 | 0.0328 | 0.0323 | 0.0335 | 0.0315 | 0.0343 | 0.0326 | 0.033577 | 0.0297 | 0.0328 | 0.0325 | 0.0322 | 0.0323 | 0.0332 | 0.0321 | 0.0329 | 0.031 | 0.0333 | 0.03777 | 0.0385 | 0.035 |  |  | 0.032725
| 1 | 5 | 0.032402 | 0.0313 | 0.032437 | 0.033478 | 0.0381 | 0.0311 | 0.0382 | 0.0328 | 0.0332 | 0.0341 | 0.0325 | 0.0346 | 0.0331 | 0.033608 | 0.0293 | 0.0316 | 0.0318 | 0.033 | 0.033 | 0.033 | 0.0329 | 0.0348 | 0.032 | 0.0333 | 0.04101 | 0.0348 | 0.0352 |  |  | 101
| 1 | 6 | 0.033482 | 0.0285 | 0.03511 | 0.033703 | 0.0346 | 0.0316 | 0.0382 | 0.0311 | 0.0307 | 0.0321 | 0.0319 | 0.0339 | 0.0334 | 0.034547 | 0.03 | 0.0321 | 0.0325 | 0.0322 | 0.0324 | 0.033 | 0.0316 | 0.0329 | 0.03 | 0.0333 | 0.04043 | 0.036 | 0.0358
| Mean |  | 0.032822 | 0.031217 | 0.032953 | 0.033206 | 0.036917 | 0.031783 | 0.038183 | 0.032117 | 0.03235 | 0.033783 | 0.032083 | 0.034033 | 0.03265 | 0.033392 | 0.0301 | 0.032133 | 0.031967 | 0.032583 | 0.03225 | 0.03275 | 0.03265 | 0.033183 | 0.030833 | 0.033167 | 0.040207 | 0.0373 | 0.034967
| Median |  | 0.032847 | 0.03165 | 0.032846 | 0.033292 | 0.0372 | 0.03135 | 0.0382 | 0.03255 | 0.03255 | 0.03385 | 0.032 | 0.0339 | 0.03255 | 0.033593 | 0.0301 | 0.03225 | 0.03215 | 0.0325 | 0.0323 | 0.03295 | 0.03275 | 0.033 | 0.031 | 0.0333 | 0.040555 | 0.03735 | 0.035
| Std Dev. |  | 0.00059805 | 0.0013717 | 0.0013605 | 0.00045507 | 0.001447 | 0.0015892 | 0.00019408 | 0.00087502 | 0.00086891 | 0.00097451 | 0.00041191 | 0.0003559 | 0.00055408 | 0.0010013 | 0.00060992 | 0.0006802 | 0.00083825 | 0.000392 | 0.00046797 | 0.00050498 | 0.00070071 | 0.0011974 | 0.00075277 | 0.00060222 | 0.0013449 | 0.0017911 | 0.00058538
| Rel.Std.Dev. |  | 1.822% | 4.394% | 4.129% | 1.37% | 3.92% | 5% | 0.5083% | 2.725% | 2.686% | 2.885% | 1.284% | 1.046% | 1.697% | 2.999% | 2.026% | 2.117% | 2.622% | 1.203% | 1.451% | 1.542% | 2.146% | 3.608% | 2.441% | 1.816% | 3.345% | 4.802% | 1.674%
| PDM3 |  | 0.2973% | -4.608% | 0.6977% | 1.47% | 12.81% | -2.876% | 16.68% | -1.858% | -1.145% | 3.235% | -1.96% | 3.999% | -0.2281% | 2.04% | -8.02% | -1.807% | -2.316% | -0.4318% | -1.45% | 0.07749% | -0.2281% | 1.402% | -5.779% | 1.351% | 22.86% | 13.98% | 6.851%
| Z-Score (Absolute) |  | 0.08 | 1.61 | 0.21 | 0.48 | 4.38 | 1.02 | 5.71 | 0.67 | 0.42 | 1.09 | 0.7 | 1.35 | 0.11 | 0.67 | 2.78 | 0.65 | 0.82 | 0.18 | 0.53 | 0 | 0.11 | 0.46 | 2.01 | 0.44 | 7.83 | 4.78 | 2.33
| Analytical results for Mo in OREAS 284 (Certified Value 2.37 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2.5978 | 2.3 | 2.1261 | < 5 | 2.65 | 2.4985 | 3 | 1.8 | 1.83 | 2.16 | 2.21 | 2.36 | 2.15 | 2.26 | 2.3564 | 2.7 | 2.61 | 2.6 | 2.42 | 2.44 | 2.44 | 3 | 2.3 |  | 2.28 | 29.1 | 3.17 | 2.4 |  | 1
| 1 | 2 | 2.5866 | 2.5 | 2.5481 | < 5 | 2.73 | 2.6026 | 3 | 1.7 | 2.03 | 2.16 | 1.97 | 2.11 | 2.21 | 2.16 | 2.5378 | 2.7 | 2.76 | 2.5 | 2.64 | 2.44 | 2.38 | 2 | 2.5 | 2.37 | 2.13 | 27.73 | 3.03 | 2.18 |  | 28
| 1 | 3 | 2.3737 | 2.5 | 2.5576 | < 5 | 2.75 | 1.6186 | 2 | 1.8 | 2.38 | 2.31 | 1.68 | 2.27 | 2.11 | 2.1 | 2.4056 | 2.4 | 2.86 | 2.5 | 2.92 | 2.33 | 2.6 | 3 | 2.5 | 2.59 | 1.98 | 26.83 | 3.1 | 2.18 |  | 16
| 1 | 4 | 2.5476 | 2.4 | 2.1417 | < 5 | 2.47 | 2.2205 | 2 | 1.7 | 2.18 | 2.16 | 2.01 | 2.12 | 2.32 | 2.01 | 2.3387 | 2.4 | 2.89 | 2.7 | 2.33 | 2.48 | 2.51 | 2 | 2.6 | 2.41 | 2.17 | 23.86 | 3.33 | 2.26 |  | 2.3721
| 1 | 5 | 2.3682 | 2.4 | 2.4732 | < 5 | 2.65 | 2.3666 | 3 | 1.7 | 2.04 | 2.17 | 2 | 2.19 | 2.18 | 2.1 | 2.5085 | 2.6 | 2.88 | 2.5 | 2.45 | 2.41 | 2.25 | 3 | 2.6 | 2.6 | 2.21 | 26.59 | 2.97 | 2.15 |  | 102
| 1 | 6 | 2.6723 | 2.3 | 2.4504 | < 5 | 2.44 | 2.4428 | 3 | 1.7 | 1.95 | 1.86 | 1.87 | 2.18 | 2.33 | 2.13 | 2.351 | 2.6 | 2.64 | 2.6 | 2.43 | 2.4 | 2.54 | 2 | 2.5 | 2.5 | 2.06 | 25.42 | 2.93 | 2.26
| Mean |  | 2.5244 | 2.4 | 2.3829 |  | 2.615 | 2.2916 | 2.6667 | 1.7333 | 2.0683 | 2.1367 | 1.9567 | 2.205 | 2.2167 | 2.1267 | 2.4163 | 2.5667 | 2.7733 | 2.5667 | 2.5317 | 2.4167 | 2.4533 | 2.5 | 2.5 | 2.494 | 2.1383 | 26.5883 | 3.0883 | 2.2383
| Median |  | 2.5671 | 2.4 | 2.4618 |  | 2.65 | 2.4047 | 3 | 1.7 | 2.035 | 2.16 | 1.985 | 2.185 | 2.195 | 2.115 | 2.381 | 2.6 | 2.81 | 2.55 | 2.44 | 2.425 | 2.475 | 2.5 | 2.5 | 2.5 | 2.15 | 26.71 | 3.065 | 2.22
| Std Dev. |  | 0.12553 | 0.089443 | 0.19729 |  | 0.13081 | 0.35384 | 0.5164 | 0.05164 | 0.19115 | 0.14787 | 0.175 | 0.095237 | 0.090259 | 0.082381 | 0.086328 | 0.13663 | 0.1242 | 0.08165 | 0.21572 | 0.050859 | 0.1258 | 0.54772 | 0.10954 | 0.10359 | 0.10722 | 1.8154 | 0.14675 | 0.091305
| Rel.Std.Dev. |  | 4.973% | 3.727% | 8.28% |  | 5.002% | 15.44% | 19.36% | 2.979% | 9.242% | 6.921% | 8.944% | 4.319% | 4.072% | 3.874% | 3.573% | 5.323% | 4.479% | 3.181% | 8.521% | 2.105% | 5.128% | 21.91% | 4.382% | 4.153% | 5.014% | 6.828% | 4.752% | 4.079%
| PDM3 |  | 6.418% | 1.175% | 0.4529% |  | 10.24% | -3.395% | 12.42% | -26.93% | -12.81% | -9.926% | -17.51% | -7.045% | -6.554% | -10.35% | 1.864% | 8.201% | 16.91% | 8.201% | 6.726% | 1.878% | 3.423% | 5.391% | 5.391% | 5.138% | -9.856% | 1021% | 30.19% | -5.64%
| Z-Score (Absolute) |  | 0.44 | 0.07 | 0.14 | 0.34 | 0.82 | 0.51 | NA | 2.81 | 1.43 | 1.15 | 1.89 | 0.87 | 0.82 | 1.19 | 0 | 0.62 | 1.47 | 0.62 | 0.47 | 0 | 0.15 | NA | 0.34 | 0.32 | 1.14 | 99.28 | 2.76 | 0.73
Results from laboratories 1.11 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 284 (Certified Value 0.159 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.16377 | 0.15 | 0.17052 | 0.186 | 0.204 | 0.2263 | 0.17 | 0.14 | 0.16 | 0.31 | 0.17 | 0.16 | 0.18 | 0.16014 | 0.15 | 0.19 | 0.13 | 0.14 | 0.158 | 0.15 | 0.13 | 0.13 | 0.12502 | 0.156 | 0.1066 | 0.177 | 0.18 |  |  | 1
| 1 | 2 | 0.15614 | 0.14 | 0.17055 | 0.18416 | 0.206 | 0.2453 | 0.17 | 0.148 | 0.16 | 0.31 | 0.17 | 0.16 | 0.16 | 0.16349 | 0.16 | 0.19 | 0.12 | 0.14 | 0.16 | 0.14 | 0.13 | 0.14 | 0.12581 | 0.156 | 0.1194 | 0.17 | 0.18 |  |  | 0
| 1 | 3 | 0.15451 | 0.14 | 0.16621 | 0.19176 | 0.198 | 0.2226 | 0.17 | 0.147 | 0.17 | 0.31 | 0.17 | 0.16 | 0.16 | 0.16586 | 0.16 | 0.18 | 0.13 | 0.14 | 0.157 | 0.14 | 0.13 | 0.14 | 0.14465 | 0.148 | 0.0999 | 0.17 | 0.18 |  |  | 16
| 1 | 4 | 0.15956 | 0.14 | 0.16172 | 0.1827 | 0.196 | 0.2109 | 0.17 | 0.15 | 0.16 | 0.31 | 0.16 | 0.16 | 0.16 | 0.16715 | 0.16 | 0.19 | 0.13 | 0.14 | 0.158 | 0.14 | 0.13 | 0.13 | 0.12252 | 0.156 | 0.0796 | 0.17 | 0.18 |  |  | 0.1593
| 1 | 5 | 0.16177 | 0.14 | 0.1731 | 0.18977 | 0.202 | 0.2202 | 0.16 | 0.149 | 0.17 | 0.31 | 0.17 | 0.16 | 0.17 | 0.17344 | 0.16 | 0.18 | 0.13 | 0.14 | 0.162 | 0.14 | 0.13 | 0.14 | 0.14628 | 0.156 | 0.099 | 0.15 | 0.19 |  |  | 103
| 1 | 6 | 0.16091 | 0.13 | 0.16696 | 0.1959 | 0.179 | 0.2289 | 0.16 | 0.137 | 0.14 | 0.26 | 0.17 | 0.16 | 0.19 | 0.17341 | 0.17 | 0.18 | 0.13 | 0.14 | 0.161 | 0.14 | 0.13 | 0.14 | 0.13031 | 0.148 | 0.1058 | 0.159 | 0.19
| Mean |  | 0.15944 | 0.14 | 0.16818 | 0.18838 | 0.1975 | 0.2257 | 0.16667 | 0.14517 | 0.16 | 0.30167 | 0.16833 | 0.16 | 0.17 | 0.16725 | 0.16 | 0.185 | 0.12833 | 0.14 | 0.15933 | 0.14167 | 0.13 | 0.13667 | 0.13243 | 0.15333 | 0.10172 | 0.166 | 0.18333
| Median |  | 0.16024 | 0.14 | 0.16874 | 0.18789 | 0.2 | 0.22445 | 0.17 | 0.1475 | 0.16 | 0.31 | 0.17 | 0.16 | 0.165 | 0.16651 | 0.16 | 0.185 | 0.13 | 0.14 | 0.159 | 0.14 | 0.13 | 0.14 | 0.12806 | 0.156 | 0.10285 | 0.17 | 0.18
| Std Dev. |  | 0.0035085 | 0.0063246 | 0.0040587 | 0.0050143 | 0.0097929 | 0.011434 | 0.005164 | 0.0053448 | 0.010954 | 0.020412 | 0.0040825 | 0 | 0.012649 | 0.0053448 | 0.0063246 | 0.0054772 | 0.0040825 | 0 | 0.0019664 | 0.0040825 | 0 | 0.005164 | 0.010417 | 0.0041312 | 0.013062 | 0.0097365 | 0.005164
| Rel.Std.Dev. |  | 2.2% | 4.518% | 2.413% | 2.662% | 4.958% | 5.066% | 3.098% | 3.682% | 6.847% | 6.767% | 2.425% | 0% | 7.441% | 3.196% | 3.953% | 2.961% | 3.181% | 0% | 1.234% | 2.882% | 0% | 3.779% | 7.866% | 2.694% | 12.84% | 5.865% | 2.817%
| PDM3 |  | 0.08915% | -12.12% | 5.571% | 18.26% | 23.98% | 41.68% | 4.624% | -8.872% | 0.439% | 89.37% | 5.67% | 0.439% | 6.716% | 4.989% | 0.439% | 16.13% | -19.44% | -12.12% | 0.02055% | -11.07% | -18.39% | -14.21% | -16.87% | -3.746% | -36.15% | 4.206% | 15.09%
| Z-Score (Absolute) |  | 0.02 | 0.74 | 0.3 | 1.04 | 1.38 | 2.42 | 0.25 | 0.55 | 0 | 5.21 | 0.31 | 0 | 0.37 | 0.27 | 0 | 0.92 | 1.16 | 0.74 | 0.02 | 0.67 | 1.1 | 0.86 | 1.01 | 0.25 | 2.14 | 0.22 | 0.86
| Analytical results for Nb in OREAS 284 (Certified Value 0.78 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [9 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [9 empty] | 2
| 1 | 1 | 1.1104 | 0.91 | 0.5 | 2.5937 | < 5 | 0.54 | 0.79 | 0.62 | 0.68 | 0.82 | 1.1499 | 1 | 1.08 | 0.7 | 2.05 | 2.18 | 0.6 | 0.24 | 0.49 | 1.83 | [9 empty] | 1
| 1 | 2 | 1.1681 | 0.83 | 0.6 | 1.4027 | < 5 | 0.6 | 0.85 | 0.66 | 0.76 | 0.8 | 1.119 | 1.4 | 1.07 | 0.7 | 2.06 | 2.18 | 0.64 | 0.24 | 0.55 | 1.66 | [9 empty] | 29
| 1 | 3 | 1.11 | 0.89 | 0.6 |  | < 5 | 0.64 | 0.74 | 0.68 | 0.82 | 0.79 | 1.2214 | 1.1 | 1.02 | 0.7 | 1.91 | 2.19 | 0.64 | 0.24 | 0.61 | 1.63 | [9 empty] | 16
| 1 | 4 | 1.0087 | 0.82 | 0.5 |  | < 5 | 0.62 | 0.98 | 0.6 | 0.92 | 0.92 | 1.1597 | 1.1 | 1.08 | 0.7 | 2.03 | 2.14 | 0.64 | 0.25 | 0.5 | 2.04 | [9 empty] | 0.78237
| 1 | 5 | 1.08 | 0.84 | 0.7 |  | < 5 | 0.83 | 0.88 | 0.65 | 0.78 | 0.82 | 1.1178 | 1 | 1.1 | 0.6 | 2.01 | 2.42 | 0.62 | 0.25 | 0.9 | 1.72 | [9 empty] | 104
| 1 | 6 | 1.1331 | 0.89 | 0.5 |  | < 5 | 0.63 | 0.87 | 0.65 | 0.79 | 0.76 | 1.2284 | 1 | 1.09 | 0.6 | 2.02 | 2.06 | 0.64 | 0.25 | 0.66 | 1.87
| Mean |  | 1.1017 | 0.86333 | 0.56667 | 1.9982 |  | 0.64333 | 0.85167 | 0.64333 | 0.79167 | 0.81833 | 1.166 | 1.1 | 1.0733 | 0.66667 | 2.0133 | 2.195 | 0.63 | 0.245 | 0.61833 | 1.7917
| Median |  | 1.1102 | 0.865 | 0.55 | 1.9982 |  | 0.625 | 0.86 | 0.65 | 0.785 | 0.81 | 1.1548 | 1.05 | 1.08 | 0.7 | 2.025 | 2.18 | 0.64 | 0.245 | 0.58 | 1.775
| Std Dev. |  | 0.054139 | 0.037771 | 0.08165 | 0.84216 |  | 0.098116 | 0.082321 | 0.028752 | 0.078592 | 0.054559 | 0.048587 | 0.15492 | 0.028048 | 0.05164 | 0.053914 | 0.12029 | 0.016733 | 0.0054772 | 0.1525 | 0.15355
| Rel.Std.Dev. |  | 4.914% | 4.375% | 14.41% | 42.15% |  | 15.25% | 9.666% | 4.469% | 9.927% | 6.667% | 4.167% | 14.08% | 2.613% | 7.746% | 2.678% | 5.48% | 2.656% | 2.236% | 24.66% | 8.57%
| PDM3 |  | 40.82% | 10.35% | -27.57% | 155.4% |  | -17.77% | 8.858% | -17.77% | 1.189% | 4.597% | 49.04% | 40.6% | 37.19% | -14.79% | 157.3% | 180.6% | -19.48% | -68.68% | -20.97% | 129%
| Z-Score (Absolute) |  | 0.67 | 0 | NA | 3.21 | 4.63 | 0.62 | 0.03 | 0.62 | 0.2 | 0.13 | 0.86 | NA | 0.59 | NA | 3.25 | 3.77 | 0.66 | 1.75 | 0.69 | 2.63
Results from laboratories 1.09, 1.20 and 1.22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 284 (Indicative Value 22.4 ppm) | [30 empty] | Indicative
Nd | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.29 |  |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*MS |  |  | 1
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.5g |  |  | 1
| 1 | 1 | 1.03: 23.7735; 1.04: 18.983; 1.05: 23.6; 1.07: 25.508; 1.09: 9.82 |  |  | 1
| 1 | 2 | 1.03: 23.4447; 1.04: 18.156; 1.05: 22.3; 1.07: 26.436; 1.09: 9.3 |  |  | 6
| 1 | 3 | 1.03: 22.6279; 1.04: 17.676; 1.05: 22.2; 1.07: 25.2369; 1.09: 9.7 |  |  | 16
| 1 | 4 | 1.03: 22.9249; 1.04: 17.154; 1.05: 21.9; 1.07: 25.5062; 1.09: 9.41 |  |  | 22.4359
| 1 | 5 | 1.03: 24.3964; 1.04: 17.462; 1.05: 23.3; 1.07: 27.2718; 1.09: 9.29 |  |  | 12
| 1 | 6 | 1.03: 23.9598; 1.04: 15.464; 1.05: 22.5; 1.07: 26.6811; 1.09: 8.84
| Mean |  | 1.03: 23.5212; 1.04: 17.4825; 1.05: 22.6333; 1.07: 26.1067; 1.09: 9.3933
| Median |  | 1.03: 23.6091; 1.04: 17.569; 1.05: 22.4; 1.07: 25.972; 1.09: 9.355
| Std Dev. |  | 1.03: 0.66051; 1.04: 1.1769; 1.05: 0.66833; 1.07: 0.80888; 1.09: 0.34697
| Rel.Std.Dev. |  | 1.03: 2.808%; 1.04: 6.732%; 1.05: 2.953%; 1.07: 3.098%; 1.09: 3.694%
| PDM3 |  | 1.03: 4.837%; 1.04: -22.08%; 1.05: 0.8799%; 1.07: 16.36%; 1.09: -58.13%
| Z-Score (Absolute) |  | 1.03: 0.17; 1.04: 1; 1.05: 0; 1.07: 0.67; 1.09: 2.57
| Analytical results for Ni in OREAS 284 (Certified Value 18.7 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 18.7906 | 19.7 | 20.3631 | 19.3515 | 19.7 | 18.906 | 20 | 17 | 19.3 | 19 | 19.2 | 18 | 19.7 | 17.8068 | 16.8 | 17.1 | 19 | 19 | 17.8 | 19 | 19 | 18 | 17.36 | 15 | 19.51 | 19.8 | 18.6 |  |  | 1
| 1 | 2 | 18.3568 | 20.1 | 19.8317 | 18.4436 | 18.9 | 20.5496 | 20 | 18 | 19 | 18.9 | 19.2 | 18.5 | 19.8 | 18.462 | 17 | 18.1 | 19 | 19 | 18.7 | 18 | 18 | 19.5 | 17.94 | 14.5 | 19.27 | 19.8 | 18.6 |  |  | 16
| 1 | 3 | 18.434 | 20 | 19.6933 | 18.5075 | 18 | 17.6834 | 20 | 18 | 19.5 | 17.2 | 19.2 | 17.8 | 18.6 | 18.7256 | 16.8 | 17.5 | 19 | 18.7 | 18.1 | 18.2 | 19 | 19.5 | 17.81 | 14.3 | 18.97 | 18.6 | 19 |  |  | 16
| 1 | 4 | 18.5697 | 19.7 | 18.7349 | 18.2949 | 18.8 | 17.1314 | 20 | 17.9 | 19.5 | 19.7 | 18.8 | 18.8 | 18 | 18.2292 | 16.6 | 17.9 | 19 | 18.6 | 18.5 | 18.2 | 18 | 19.1 | 17.57 | 14.5 | 17.86 | 18.1 | 19 |  |  | 18.6571
| 1 | 5 | 18.5583 | 19.8 | 19.9428 | 18.6956 | 19.2 | 17.6097 | 20 | 17.9 | 19 | 19.8 | 19.5 | 18.7 | 18.6 | 17.8431 | 16.7 | 17.4 | 19 | 18.6 | 18.5 | 18.6 | 19 | 19.7 | 18.12 | 14.4 | 19.33 | 17.4 | 18.9 |  |  | 105
| 1 | 6 | 18.7751 | 20.2 | 20.5518 | 18.6468 | 18.6 | 17.7395 | 20 | 16.9 | 18.6 | 17.3 | 19.6 | 18.5 | 18.7 | 18.4169 | 17.3 | 17.6 | 19 | 18.4 | 18.6 | 18.3 | 18 | 18.5 | 17.2 | 14.2 | 19.22 | 19.2 | 19
| Mean |  | 18.5808 | 19.9167 | 19.8529 | 18.6566 | 18.8667 | 18.2699 | 20 | 17.6167 | 19.15 | 18.65 | 19.25 | 18.3833 | 18.9 | 18.2473 | 16.8667 | 17.6 | 19 | 18.7167 | 18.3667 | 18.3833 | 18.5 | 19.05 | 17.6667 | 14.4833 | 19.0267 | 18.8167 | 18.85
| Median |  | 18.564 | 19.9 | 19.8873 | 18.5772 | 18.85 | 17.7114 | 20 | 17.9 | 19.15 | 18.95 | 19.2 | 18.5 | 18.65 | 18.3231 | 16.8 | 17.55 | 19 | 18.65 | 18.5 | 18.25 | 18.5 | 19.3 | 17.69 | 14.45 | 19.245 | 18.9 | 18.95
| Std Dev. |  | 0.17557 | 0.2137 | 0.6376 | 0.36958 | 0.57155 | 1.2617 | 0 | 0.51929 | 0.35071 | 1.1432 | 0.28107 | 0.39707 | 0.70427 | 0.36368 | 0.25033 | 0.35777 | 0 | 0.24014 | 0.34448 | 0.36009 | 0.54772 | 0.66858 | 0.35268 | 0.27869 | 0.59775 | 0.96419 | 0.19748
| Rel.Std.Dev. |  | 0.9449% | 1.073% | 3.212% | 1.981% | 3.029% | 6.906% | 0% | 2.948% | 1.831% | 6.13% | 1.46% | 2.16% | 3.726% | 1.993% | 1.484% | 2.033% | 0% | 1.283% | 1.876% | 1.959% | 2.961% | 3.51% | 1.996% | 1.924% | 3.142% | 5.124% | 1.048%
| PDM3 |  | -0.4091% | 6.751% | 6.41% | -0.002359% | 1.123% | -2.075% | 7.198% | -5.577% | 2.642% | -0.038% | 3.178% | -1.467% | 1.302% | -2.197% | -9.596% | -5.666% | 1.838% | 0.3193% | -1.557% | -1.467% | -0.842% | 2.106% | -5.309% | -22.37% | 1.981% | 0.8553% | 1.034%
| Z-Score (Absolute) |  | 0.13 | 2.25 | 2.14 | 0.01 | 0.38 | 0.68 | NA | 1.85 | 0.89 | 0.01 | 1.06 | 0.48 | 0.44 | 0.72 | 3.18 | 1.88 | NA | 0.11 | 0.51 | 0.48 | NA | 0.71 | 1.76 | 7.43 | 0.67 | 0.29 | 0.35
Results from laboratories 1.11, 1.22 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 284 (Certified Value 0.064 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.065203 | 0.0639 | 0.065286 | 0.062696 | 0.063 | 0.0629 | 0.0578 | 0.065 | 0.066 | 0.062 | 0.064 | 0.063 | 0.062203 | 0.07 | 0.0622 | 0.0665 | 0.0688 | 0.061 | 0.06 | 0.0618 | 0.0601 | 0.0676 | 0.07433 | 0.0698 | 0.0662 |  |  |  |  | 1
| 1 | 2 | 0.062674 | 0.0628 | 0.064966 | 0.060021 | 0.062 | 0.0631 | 0.0611 | 0.066 | 0.066 | 0.062 | 0.064 | 0.062 | 0.062628 | 0.069 | 0.063 | 0.0651 | 0.0698 | 0.064 | 0.06 | 0.0612 | 0.0647 | 0.0659 | 0.07376 | 0.065 | 0.0675 |  |  |  |  | 17
| 1 | 3 | 0.062584 | 0.0633 | 0.063714 | 0.0613 | 0.061 | 0.063 | 0.0608 | 0.066 | 0.065 | 0.062 | 0.064 | 0.061 | 0.061396 | 0.069 | 0.062 | 0.0659 | 0.0684 | 0.061 | 0.06 | 0.0623 | 0.0651 | 0.0663 | 0.07249 | 0.064 | 0.0689 |  |  |  |  | 16
| 1 | 4 | 0.063819 | 0.0643 | 0.062258 | 0.061361 | 0.062 | 0.0631 | 0.061 | 0.065 | 0.065 | 0.061 | 0.065 | 0.063 | 0.063701 | 0.068 | 0.0638 | 0.0664 | 0.069 | 0.062 | 0.06 | 0.0602 | 0.0646 | 0.0668 | 0.06703 | 0.066 | 0.0688 |  |  |  |  | 0.064041
| 1 | 5 | 0.063566 | 0.0621 | 0.06468 | 0.062201 | 0.062 | 0.0628 | 0.0604 | 0.066 | 0.066 | 0.063 | 0.066 | 0.064 | 0.06444 | 0.068 | 0.063 | 0.0658 | 0.0698 | 0.063 | 0.06 | 0.0608 | 0.068 | 0.0668 | 0.07351 | 0.069 | 0.0687 |  |  |  |  | 106
| 1 | 6 | 0.063687 | 0.0638 | 0.067453 | 0.061864 | 0.062 | 0.0627 | 0.0575 | 0.062 | 0.061 | 0.062 | 0.064 | 0.065 | 0.061054 | 0.068 | 0.0631 | 0.0658 | 0.0679 | 0.061 | 0.06 | 0.0602 | 0.0632 | 0.0659 | 0.07244 | 0.0685 | 0.0703
| Mean |  | 0.063589 | 0.063367 | 0.064726 | 0.061574 | 0.062 | 0.062933 | 0.059767 | 0.065 | 0.064833 | 0.062 | 0.0645 | 0.063 | 0.06257 | 0.068667 | 0.06285 | 0.065917 | 0.06895 | 0.062 | 0.06 | 0.061083 | 0.064283 | 0.06655 | 0.07226 | 0.06705 | 0.0684
| Median |  | 0.063627 | 0.06355 | 0.064823 | 0.061612 | 0.062 | 0.06295 | 0.0606 | 0.0655 | 0.0655 | 0.062 | 0.064 | 0.063 | 0.062416 | 0.0685 | 0.063 | 0.06585 | 0.0689 | 0.0615 | 0.06 | 0.061 | 0.06465 | 0.06655 | 0.073 | 0.06725 | 0.06875
| Std Dev. |  | 0.00095042 | 0.00080911 | 0.0017277 | 0.00092382 | 0.00063246 | 0.0001633 | 0.0016597 | 0.0015492 | 0.0019408 | 0.00063246 | 0.00083666 | 0.0014142 | 0.0013107 | 0.0008165 | 0.00065651 | 0.00050365 | 0.00075829 | 0.0012649 | 0 | 0.00085421 | 0.0025857 | 0.00065345 | 0.0026659 | 0.0023696 | 0.0013971
| Rel.Std.Dev. |  | 1.495% | 1.277% | 2.669% | 1.5% | 1.02% | 0.2595% | 2.777% | 2.383% | 2.994% | 1.02% | 1.297% | 2.245% | 2.095% | 1.189% | 1.045% | 0.7641% | 1.1% | 2.04% | 0% | 1.398% | 4.022% | 0.9819% | 3.689% | 3.534% | 2.043%
| PDM3 |  | -0.7054% | -1.052% | 1.07% | -3.852% | -3.186% | -1.729% | -6.674% | 1.498% | 1.238% | -3.186% | 0.7174% | -1.625% | -2.296% | 7.224% | -1.859% | 2.93% | 7.666% | -3.186% | -6.309% | -4.618% | 0.3791% | 3.919% | 12.83% | 4.699% | 6.807%
| Z-Score (Absolute) |  | 0 | 0.09 | 0.48 | 0.86 | 0.67 | 0.28 | 1.62 | 0.6 | 0.53 | 0.67 | 0.39 | 0.25 | 0.43 | 2.16 | 0.31 | 0.99 | 2.28 | 0.67 | 1.52 | 1.06 | 0.29 | 1.26 | 3.68 | 1.47 | 2.04
| Analytical results for Pb in OREAS 284 (Certified Value 4.79 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 5.3751 | 5.1 | 5.3307 | < 5 | 5.9 |  | < 5 | 6 | 3.85 | 4.9 | 5 | 4.8 | 4.8 | 5.6 | 4.5254 | 4.4 | 4.4 | 4.9 | 6.1 | 4.44 | 4.8 | 5 | 4.2 | 4.48 | 50.648 | 9.3 | 4.1 |  |  | 1
| 1 | 2 | 5.7952 | 5 | 5.5161 | < 5 | 5.6 | 7.1342 | < 5 | 12 | 4.09 | 4.9 | 5.1 | 4.7 | 4.6 | 5 | 4.5524 | 4.6 | 4.4 | 4.9 | 5.3 | 4.56 | 5.3 | 5 | 4.6 | 4.15 | 51.336 | 7.4 | 4.1 |  |  | 18
| 1 | 3 | 5.6565 | 5 | 5.5847 | < 5 | 5.7 | 8.9926 | 5 | 9 | 4.33 | 4.9 | 4.6 | 4.8 | 4.5 | 4.8 | 4.849 | 5.6 | 4.4 | 5.1 | 5.2 | 4.43 | 4.7 | 5 | 4.6 | 4.18 | 50.888 | 7.1 | 4 |  |  | 16
| 1 | 4 | 5.6873 | 5 | 5.356 | < 5 | 5.8 | 4.5065 | < 5 | 9 | 4.16 | 4.9 | 5.3 | 4.6 | 4.9 | 4.7 | 4.6951 | 4.5 | 4.4 | 5 | 5.1 | 4.81 | 4.9 | 5 | 4.4 | 4.16 | 46.632 | 8 | 3.9 |  |  | 4.794
| 1 | 5 | 5.8154 | 4.9 | 5.6505 | < 5 | 5.9 |  | < 5 | 6 | 4.05 | 5 | 4.9 | 4.8 | 4.7 | 4.9 | 4.9227 | 4.5 | 4.2 | 5.1 | 4.8 | 4.77 | 4.9 | 5 | 4.7 | 4.06 | 52.4 | 7.2 | 3.8 |  |  | 107
| 1 | 6 | 5.9612 | 5.1 | 5.4687 | < 5 | 5.8 | 6.7557 | < 5 | 5 | 4.24 | 4.1 | 4.6 | 4.7 | 4.8 | 4.6 | 4.5108 | 4.6 | 4.2 | 4.9 | 4.9 | 4.35 | 4.7 | 5 | 4.5 | 3.92 | 51.728 | 7.3 | 4.1
| Mean |  | 5.7151 | 5.0167 | 5.4844 |  | 5.7833 | 6.8472 | 5 | 7.8333 | 4.12 | 4.7833 | 4.9167 | 4.7333 | 4.7167 | 4.9333 | 4.6759 | 4.7 | 4.3333 | 4.9833 | 5.2333 | 4.56 | 4.8833 | 5 | 4.5 | 4.1583 | 50.6053 | 7.7167 | 4
| Median |  | 5.7413 | 5 | 5.4924 |  | 5.8 | 6.945 | 5 | 7.5 | 4.125 | 4.9 | 4.95 | 4.75 | 4.75 | 4.85 | 4.6238 | 4.55 | 4.4 | 4.95 | 5.15 | 4.5 | 4.85 | 5 | 4.55 | 4.155 | 51.112 | 7.35 | 4.05
| Std Dev. |  | 0.19842 | 0.075277 | 0.1257 |  | 0.1169 | 1.8414 |  | 2.6394 | 0.16661 | 0.33714 | 0.27869 | 0.08165 | 0.1472 | 0.3559 | 0.17688 | 0.44721 | 0.10328 | 0.098319 | 0.46332 | 0.19079 | 0.22286 | 0 | 0.17889 | 0.18465 | 2.0437 | 0.83766 | 0.12649
| Rel.Std.Dev. |  | 3.472% | 1.501% | 2.292% |  | 2.021% | 26.89% |  | 33.7% | 4.044% | 7.048% | 5.668% | 1.725% | 3.121% | 7.214% | 3.783% | 9.515% | 2.383% | 1.973% | 8.853% | 4.184% | 4.564% | 0% | 3.975% | 4.441% | 4.039% | 10.86% | 3.162%
| PDM3 |  | 19.21% | 4.644% | 14.4% |  | 20.64% | 42.83% | 4.297% | 63.4% | -14.06% | -0.223% | 2.558% | -1.266% | -1.614% | 2.906% | -2.464% | -1.961% | -9.61% | 3.949% | 9.164% | -4.882% | 1.863% | 4.297% | -6.133% | -13.26% | 955.6% | 60.96% | -16.56%
| Z-Score (Absolute) |  | 1.62 | 0.34 | 1.2 | 4.29 | 1.75 | 3.7 | NA | NA | 1.31 | 0.09 | 0.15 | 0.18 | 0.21 | 0.18 | 0.29 | 0.25 | 0.92 | 0.28 | 0.74 | 0.5 | 0.09 | NA | 0.61 | 1.24 | 84.18 | 5.3 | 1.53
Results from laboratories 1.11, 1.12 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 284 (Indicative Value 19.5 ppb) | [30 empty] | Indicative
Pd | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab |  |  | 1
| No | No | 1.03: 1.04; 1.04: 1.19; 1.05: 1.23 |  |  | ppb
|  |  | 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*OES/MS |  |  | 1
|  |  | 1.03: 01g; 1.04: 01g; 1.05: 0.5g |  |  | 1
| 1 | 1 | 1.03: < 10; 1.04: < 10; 1.05: 21 |  |  | 1
| 1 | 2 | 1.03: < 10; 1.04: < 10; 1.05: 19 |  |  | 7
| 1 | 3 | 1.03: < 10; 1.04: < 10; 1.05: 18 |  |  | 16
| 1 | 4 | 1.03: < 10; 1.04: < 10; 1.05: 20 |  |  | 19.5
| 1 | 5 | 1.03: < 10; 1.04: < 10; 1.05: 20 |  |  | 13
| 1 | 6 | 1.03: < 10; 1.04: < 10; 1.05: 19
| Mean |  | 1.05: 19.5
| Median |  | 1.05: 19.5
| Std Dev. |  | 1.05: 1.0488
| Rel.Std.Dev. |  | 1.05: 5.379%
PDM3 | bar: 1.05
Z-Score (Absolute) | bar: 1.03 to 1.05
| Analytical results for Pr in OREAS 284 (Indicative Value 5.51 ppm) | [30 empty] | Indicative
Pr | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.29 |  |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*MS |  |  | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.5g |  |  | 2
| 1 | 1 | 1.03: 5.8648; 1.04: 4.743; 1.05: 5.6; 1.07: 6.3368; 1.09: 2.11 |  |  | 1
| 1 | 2 | 1.03: 5.8471; 1.04: 4.55; 1.05: 5.5; 1.07: 6.4201; 1.09: 1.96 |  |  | 4
| 1 | 3 | 1.03: 5.6664; 1.04: 4.501; 1.05: 5.4; 1.07: 6.2604; 1.09: 2.07 |  |  | 16
| 1 | 4 | 1.03: 5.6061; 1.04: 4.304; 1.05: 5.3; 1.07: 5.982; 1.09: 2 |  |  | 5.5062
| 1 | 5 | 1.03: 5.9994; 1.04: 4.372; 1.05: 5.7; 1.07: 6.6237; 1.09: 1.98 |  |  | 10
| 1 | 6 | 1.03: 5.9889; 1.04: 3.865; 1.05: 5.5; 1.07: 6.2174; 1.09: 1.85
| Mean |  | 1.03: 5.8288; 1.04: 4.3892; 1.05: 5.5; 1.07: 6.3067; 1.09: 1.995
| Median |  | 1.03: 5.856; 1.04: 4.4365; 1.05: 5.5; 1.07: 6.2986; 1.09: 1.99
| Std Dev. |  | 1.03: 0.1627; 1.04: 0.29862; 1.05: 0.14142; 1.07: 0.21438; 1.09: 0.09094
| Rel.Std.Dev. |  | 1.03: 2.791%; 1.04: 6.804%; 1.05: 2.571%; 1.07: 3.399%; 1.09: 4.558%
| PDM3 |  | 1.03: 5.859%; 1.04: -20.29%; 1.05: -0.1121%; 1.07: 14.54%; 1.09: -63.77%
| Z-Score (Absolute) |  | 1.03: 0.27; 1.04: 0.93; 1.05: 0; 1.07: 0.67; 1.09: 2.93
| Analytical results for Pt in OREAS 284 (Indicative Value < 5 ppb) | [30 empty] | Indicative
Pt | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab |  |  | 1
| No | No | 1.03: 1.04; 1.04: 1.19; 1.05: 1.23 |  |  | ppb
|  |  | 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*OES/MS |  |  | 2
|  |  | 1.03: 01g; 1.04: 01g; 1.05: 0.5g |  |  | 2
| 1 | 1 | 1.03: < 5; 1.04: < 5; 1.05: < 10 |  |  | 1
| 1 | 2 | 1.03: < 5; 1.04: < 5; 1.05: < 10 |  |  | 5
| 1 | 3 | 1.03: < 5; 1.04: < 5; 1.05: < 10 |  |  | 16
| 1 | 4 | 1.03: < 5; 1.04: < 5; 1.05: < 10 |  |  | < 5
| 1 | 5 | 1.03: < 5; 1.04: < 5; 1.05: < 10 |  |  | 11
| 1 | 6 | 1.03: < 5; 1.04: < 5; 1.05: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.03 to 1.05
| Analytical results for Rb in OREAS 284 (Certified Value 105 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 0
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [11 empty] | 0
| 1 | 1 | 102.4151 | 110.7 | 94.2 | 102 | 116 | 108 | 111 | 104 | 92.5725 | 101.1 | 103.2 | 113 | 100 | 107 | 103.43 | 107 | 127 | 106.7 | [11 empty] | 1
| 1 | 2 | 104.8452 | 111.01 | 89.6 | 105.5 | 114.5 | 106 | 111 | 105 | 97.1785 | 101.5 | 105.9 | 110 | 101 | 104 | 109.62 | 103 | 125 | 110.1 | [11 empty] | 34
| 1 | 3 | 102.8323 | 110.51 | 91.5 | 104 | 107 | 105.5 | 111 | 101.5 | 93.0367 | 101.1 | 105.6 | 113 | 100 | 104 | 112.16 | 104 | 125 | 110.8 | [11 empty] | 16
| 1 | 4 | 102.2063 | 109.8 | 92.1 | 105 | 115 | 104 | 113.5 | 98.1 | 97.0122 | 100 | 106.6 | 112 | 100 | 102 | 109.78 | 105 | 126 | 110.5 | [11 empty] | 105.167
| 1 | 5 | 104.1347 | 110.47 | 96.1 | 105.5 | 113 | 110 | 116.5 | 101.5 | 96.9471 | 100.6 | 105.2 | 113 | 101 | 107 | 115.2 | 103 | 121 | 110 | [11 empty] | 108
| 1 | 6 | 105.212 | 111.28 | 87.1 | 103.5 | 109 | 105.5 | 112.5 | 101 | 97.5645 | 101.7 | 107.2 | 113 | 101 | 107 | 110.12 | 100 | 123 | 113
| Mean |  | 103.6076 | 110.6283 | 91.7667 | 104.25 | 112.4167 | 106.5 | 112.5833 | 101.85 | 95.7186 | 101 | 105.6167 | 112.3333 | 100.5 | 105.1667 | 110.0517 | 103.6667 | 124.5 | 110.1833
| Median |  | 103.4835 | 110.605 | 91.8 | 104.5 | 113.75 | 105.75 | 111.75 | 101.5 | 96.9796 | 101.1 | 105.75 | 113 | 100.5 | 105.5 | 109.95 | 103.5 | 125 | 110.3
| Std Dev. |  | 1.2939 | 0.51004 | 3.2048 | 1.3693 | 3.6113 | 2.1448 | 2.1775 | 2.4321 | 2.2721 | 0.61968 | 1.3834 | 1.2111 | 0.54772 | 2.137 | 3.8746 | 2.3381 | 2.1679 | 2.0292
| Rel.Std.Dev. |  | 1.249% | 0.461% | 3.492% | 1.313% | 3.212% | 2.014% | 1.934% | 2.388% | 2.374% | 0.6135% | 1.31% | 1.078% | 0.545% | 2.032% | 3.521% | 2.255% | 1.741% | 1.842%
| PDM3 |  | -1.483% | 5.193% | -12.74% | -0.872% | 6.893% | 1.267% | 7.052% | -3.154% | -8.984% | -3.962% | 0.4275% | 6.814% | -4.438% | -0.0003452% | 4.645% | -1.427% | 18.38% | 4.77%
| Z-Score (Absolute) |  | 0.25 | 0.75 | 1.94 | 0.16 | 1 | 0.16 | 1.03 | 0.51 | 1.38 | 0.63 | 0.03 | 0.99 | 0.7 | 0.03 | 0.66 | 0.25 | 2.73 | 0.68
| Analytical results for Re in OREAS 284 (Indicative Value 0.001 ppm) | [30 empty] | Indicative
Re | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab; 1.17: Lab; 1.18: Lab; 1.19: Lab; 1.20: Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.27 | 1.31 | 1.32 | [13 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [13 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.25g | [13 empty] | 3
| 1 | 1 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.005 | 0.002 | < 0.05 | 0.001 |  | 0.0015 | < 0.1 | [13 empty] | 1
| 1 | 2 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.001 | < 0.05 | < 0.005 | 0.002 | < 0.05 | 0.001 |  | 0.001 | < 0.1 | [13 empty] | 2
| 1 | 3 | < 0.01 | < 0.001 | 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.005 | < 0.002 | < 0.05 | 0.002 |  | 0.001 | < 0.1 | [13 empty] | 16
| 1 | 4 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.001 | < 0.05 | < 0.005 | < 0.002 | < 0.05 | 0.002 | 0.001 | 0.001 | < 0.1 | [13 empty] | 0.0011909
| 1 | 5 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.001 | 0.001 | < 0.05 | < 0.005 | < 0.002 | < 0.05 | 0.001 |  | 0.001 | < 0.1 | [13 empty] | 12
| 1 | 6 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.001 | < 0.05 | < 0.005 | < 0.002 | < 0.05 | 0.002 |  | 0.001 | < 0.1
| Mean |  |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.0016 | 0.001 |  |  | 0.002 |  | 0.0015 | 0.001 | 0.0010833
| Median |  |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.001 |  |  | 0.002 |  | 0.0015 | 0.001 | 0.001
| Std Dev. |  | 1.09: 0; 1.14: 0; 1.15: 0; 1.16: 0; 1.17: 0.00054772; 1.18: 0; 1.21: 0; 1.23: 0.00054772; 1.31: 0.00020412
| Rel.Std.Dev. |  | 1.09: 0%; 1.14: 0%; 1.15: 0%; 1.16: 0%; 1.17: 34.23%; 1.18: 0%; 1.21: 0%; 1.23: 36.51%; 1.31: 18.84%
| PDM3 |  |  | -16.03% | -16.03% | -16.03% | -16.03% | -16.03% | 34.35% | -16.03% |  |  | 67.94% |  | 25.95% | -16.03% | -9.033%
| Z-Score (Absolute) |  | 10.23 | 1.69 | 0.56 | 0.79 | 0.56 | 0.56 | 0.56 | 0.56 | 64.17 | 3.48 | 0.34 | 64.17 | 0.79 | 0.56 | 0.34 | 131.6
| Analytical results for S in OREAS 284 (Certified Value 0.255 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.23941 | 0.28162 | 0.2567 | 0.27233 | 0.25913 | 0.252 | 0.2226 | 0.3 | 0.2291 | 0.24 | 0.27 | 0.29 | 0.26 | 0.28 | 0.29 | 0.25082 | 0.25 | 0.27 | 0.23 | 0.26 | 0.24 | 0.27 | 0.2505 | 0.22 | 0.266 | 0.278 | 0.27 |  |  | 1
| 1 | 2 | 0.23205 | 0.26416 | 0.2545 | 0.25407 | 0.25701 | 0.247 | 0.2363 | 0.3 | 0.2308 | 0.26 | 0.27 | 0.29 | 0.26 | 0.27 | 0.3 | 0.24422 | 0.24 | 0.27 | 0.24 | 0.26 | 0.24 | 0.25 | 0.247 | 0.23 | 0.261 | 0.27 | 0.27 |  |  | 11
| 1 | 3 | 0.23638 | 0.26359 | 0.2583 | 0.2595 | 0.25467 | 0.24 | 0.2168 | 0.29 | 0.2344 | 0.26 | 0.28 | 0.28 | 0.26 | 0.27 | 0.29 | 0.2419 | 0.25 | 0.26 | 0.25 | 0.25 | 0.26 | 0.26 | 0.2541 | 0.22 | 0.262 | 0.27 | 0.28 |  |  | 16
| 1 | 4 | 0.23099 | 0.2684 | 0.2607 | 0.25208 | 0.255 | 0.243 | 0.2078 | 0.3 | 0.2319 | 0.26 | 0.27 | 0.28 | 0.26 | 0.28 | 0.3 | 0.26856 | 0.25 | 0.27 | 0.25 | 0.25 | 0.26 | 0.26 | 0.2403 | 0.22 | 0.269 | 0.25 | 0.28 |  |  | 0.255
| 1 | 5 | 0.23336 | 0.26501 | 0.2532 | 0.25843 | 0.25193 | 0.243 | 0.2093 | 0.3 | 0.2339 | 0.26 | 0.28 | 0.29 | 0.26 | 0.28 | 0.3 | 0.27012 | 0.25 | 0.27 | 0.25 | 0.26 | 0.26 | 0.26 | 0.247 | 0.25 | 0.264 | 0.26 | 0.27 |  |  | 109
| 1 | 6 | 0.23267 | 0.27385 | 0.2573 | 0.25507 | 0.25163 | 0.247 | 0.2219 | 0.3 | 0.2305 | 0.25 | 0.25 | 0.28 | 0.26 | 0.27 | 0.31 | 0.2549 | 0.26 | 0.27 | 0.25 | 0.25 | 0.27 | 0.26 | 0.237 | 0.22 | 0.262 | 0.275 | 0.28
| Mean |  | 0.23414 | 0.26944 | 0.25678 | 0.25858 | 0.2549 | 0.24533 | 0.21912 | 0.29833 | 0.23177 | 0.255 | 0.27 | 0.285 | 0.26 | 0.275 | 0.29833 | 0.25509 | 0.25 | 0.26833 | 0.245 | 0.255 | 0.255 | 0.26 | 0.24598 | 0.22667 | 0.264 | 0.26717 | 0.275
| Median |  | 0.23301 | 0.26671 | 0.257 | 0.25675 | 0.25484 | 0.245 | 0.21935 | 0.3 | 0.23135 | 0.26 | 0.27 | 0.285 | 0.26 | 0.275 | 0.3 | 0.25286 | 0.25 | 0.27 | 0.25 | 0.255 | 0.26 | 0.26 | 0.247 | 0.22 | 0.263 | 0.27 | 0.275
| Std Dev. |  | 0.0031575 | 0.0070779 | 0.0026821 | 0.0072765 | 0.002896 | 0.0042269 | 0.01044 | 0.0040825 | 0.0020569 | 0.0083666 | 0.010954 | 0.0054772 | 0 | 0.0054772 | 0.0075277 | 0.011978 | 0.0063246 | 0.0040825 | 0.0083666 | 0.0054772 | 0.012247 | 0.0063246 | 0.0063458 | 0.012111 | 0.0030332 | 0.0104 | 0.0054772
| Rel.Std.Dev. |  | 1.349% | 2.627% | 1.044% | 2.814% | 1.136% | 1.723% | 4.764% | 1.368% | 0.8875% | 3.281% | 4.057% | 1.922% | 0% | 1.992% | 2.523% | 4.696% | 2.53% | 1.521% | 3.415% | 2.148% | 4.803% | 2.433% | 2.58% | 5.343% | 1.149% | 3.893% | 1.992%
| PDM3 |  | -8.177% | 5.664% | 0.7013% | 1.406% | -0.03852% | -3.789% | -14.07% | 17% | -9.109% | 0.00191% | 5.884% | 11.77% | 1.963% | 7.845% | 17% | 0.03556% | -1.959% | 5.231% | -3.92% | 0.00191% | 0.00191% | 1.963% | -3.534% | -11.11% | 3.531% | 4.773% | 7.845%
| Z-Score (Absolute) |  | 1.33 | 0.75 | 0 | 0.11 | 0.11 | 0.67 | 2.22 | 2.45 | 1.47 | 0.11 | 0.78 | 1.66 | 0.19 | 1.07 | 2.45 | 0.1 | 0.4 | 0.68 | 0.69 | 0.11 | 0.11 | 0.19 | 0.64 | 1.77 | 0.43 | 0.61 | 1.07
| Analytical results for Sb in OREAS 284 (Certified Value 0.25 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 0.29166 | 0.3 | 6.19 | < 5 | 0.27 | < 5 | 0.27 | 0.28 | 0.26 | 0.26 | 0.29 | 0.25 | 0.35 | 0.24713 | 0.17 | 0.17 | 0.14 | 0.28 | 0.24 | 0.26 | < 2 | 0.17 | 0.32 | 0.657 | 0.51 | 0.13 |  |  |  | 1
| 1 | 2 | 0.29775 | 0.28 | 3.9428 | < 5 | 0.26 | < 5 | 0.28 | 0.26 | 0.25 | 0.25 | 0.26 | 0.25 | 0.28 | 0.24802 | 0.17 | 0.18 | 0.16 | 0.29 | 0.23 | 0.24 | < 2 | 0.18 | 0.25 | 0.576 | 0.39 | 0.16 |  |  |  | 36
| 1 | 3 | 0.27742 | 0.3 | 4.1073 | < 5 | 0.26 | < 5 | 0.27 | 0.28 | 0.24 | 0.21 | 0.28 | 0.24 | 0.25 | 0.28546 | 0.17 | 0.17 | 0.16 | 0.27 | 0.22 | 0.25 | < 2 | 0.2 | 0.26 | 0.585 | 0.46 | 0.17 |  |  |  | 16
| 1 | 4 | 0.2787 | 0.29 | 0.8323 | < 5 | 0.28 | < 5 | 0.27 | 0.28 | 0.25 | 0.27 | 0.25 | 0.28 | 0.23 | 0.2902 | 0.17 | 0.18 | 0.14 | 0.28 | 0.21 | 0.24 | < 2 | 0.19 | 0.25 | 0.513 | 0.33 | 0.13 |  |  |  | 0.24621
| 1 | 5 | 0.27984 | 0.28 | 4.2941 | < 5 | 0.27 | < 5 | 0.26 | 0.27 | 0.25 | 0.26 | 0.27 | 0.25 | 0.26 | 0.24713 | 0.16 | 0.17 | 0.15 | 0.27 | 0.21 | 0.27 | < 2 | 0.24 | 0.26 | 0.342 | 0.34 | 0.12 |  |  |  | 110
| 1 | 6 | 0.27447 | 0.25 | 3.26 | < 5 | 0.27 | < 5 | 0.26 | 0.26 | 0.22 | 0.24 | 0.26 | 0.26 | 0.24 | 0.23469 | 0.18 | 0.17 | 0.15 | 0.28 | 0.23 | 0.23 | < 2 | 0.18 | 0.24 | 0.549 | 0.29 | 0.13
| Mean |  | 0.28331 | 0.28333 | 3.7711 |  | 0.26833 |  | 0.26833 | 0.27167 | 0.245 | 0.24833 | 0.26833 | 0.255 | 0.26833 | 0.25877 | 0.17 | 0.17333 | 0.15 | 0.27833 | 0.22333 | 0.24833 |  | 0.19333 | 0.26333 | 0.537 | 0.38667 | 0.14
| Median |  | 0.27927 | 0.285 | 4.0251 |  | 0.27 |  | 0.27 | 0.275 | 0.25 | 0.255 | 0.265 | 0.25 | 0.255 | 0.24758 | 0.17 | 0.17 | 0.15 | 0.28 | 0.225 | 0.245 |  | 0.185 | 0.255 | 0.5625 | 0.365 | 0.13
| Std Dev. |  | 0.0092147 | 0.018619 | 1.7415 |  | 0.0075277 |  | 0.0075277 | 0.0098319 | 0.013784 | 0.02137 | 0.01472 | 0.013784 | 0.043551 | 0.023095 | 0.0063246 | 0.005164 | 0.0089443 | 0.0075277 | 0.012111 | 0.01472 |  | 0.025033 | 0.028752 | 0.10674 | 0.084063 | 0.02
| Rel.Std.Dev. |  | 3.253% | 6.571% | 46.18% |  | 2.805% |  | 2.805% | 3.619% | 5.626% | 8.605% | 5.486% | 5.406% | 16.23% | 8.925% | 3.72% | 2.979% | 5.963% | 2.705% | 5.423% | 5.927% |  | 12.95% | 10.92% | 19.88% | 21.74% | 14.29%
| PDM3 |  | 15.07% | 15.08% | 1432% |  | 8.986% |  | 8.986% | 10.34% | -0.4907% | 0.8631% | 8.986% | 3.571% | 8.986% | 5.104% | -30.95% | -29.6% | -39.08% | 13.05% | -9.291% | 0.8631% |  | -21.48% | 6.956% | 118.1% | 57.05% | -43.14%
| Z-Score (Absolute) |  | 0.47 | 0.47 | 109.01 | 69.45 | 0 | 69.45 | 0 | 0.1 | 0.73 | 0.62 | 0 | 0.41 | 0 | 0.3 | 3.06 | 2.96 | 3.68 | 0.31 | 1.4 | 0.62 | 22.77 | 2.33 | 0.16 | 8.36 | 3.68 | 3.99
| Analytical results for Sc in OREAS 284 (Certified Value 8.16 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 8.3648 | 8.2 | 8.5 | 6.3378 | 8 | 7.6 | 8.1 | 8.7 | 8.2 | 8.8 | 8.2 | 8.4976 | 8.3 | 8.4 | 7 | 8.1 | 8 | 8.2 | 8 | 7.8 | 5.2 | 8.45 | 11.1 | 7.9 | [5 empty] | 1
| 1 | 2 | 8.1334 | 8.1 | 8.1 | 7.288 | 8 | 7.9 | 8 | 8.4 | 8.2 | 9.1 | 8.4 | 8.3972 | 8.1 | 8.6 | 7 | 8.3 | 7.7 | 7.9 | 8 | 8.1 | 5.1 | 8.37 | 10.6 | 8 | [5 empty] | 37
| 1 | 3 | 8.149 | 8.1 | 8 | 7.6763 | 8 | 7.8 | 8.2 | 7.8 | 8 | 8.6 | 8.1 | 8.5999 | 8 | 8.3 | 7 | 8 | 7.9 | 8.1 | 8 | 8.4 | 5.1 | 8.22 | 10.6 | 8.2 | [5 empty] | 16
| 1 | 4 | 8.2847 | 8.3 | 7.9 | 6.97 | 8 | 7.9 | 8.1 | 9.4 | 8 | 9.2 | 7.9 | 8.5327 | 7.9 | 8.5 | 8 | 8.2 | 7.8 | 8.1 | 8 | 7.9 | 5.1 | 7.88 | 11.1 | 7.9 | [5 empty] | 8.1571
| 1 | 5 | 8.1966 | 8 | 8.5 | 7.6993 | 8 | 7.9 | 8 | 8.9 | 8.4 | 8.9 | 8.2 | 8.6263 | 7.9 | 8.6 | 7 | 8.2 | 7.9 | 8.3 | 8 | 8.4 | 5.1 | 8.49 | 9.5 | 8 | [5 empty] | 111
| 1 | 6 | 8.3874 | 7.3 | 7.4 | 7.3833 | 8 | 7.6 | 9.5 | 9.4 | 8.2 | 8.8 | 8.1 | 8.4212 | 7.9 | 8.4 | 7 | 8.1 | 8.1 | 8.1 | 8 | 8.2 | 5 | 8.36 | 10 | 8.4
| Mean |  | 8.2527 | 8 | 8.0667 | 7.2258 | 8 | 7.7833 | 8.3167 | 8.7667 | 8.1667 | 8.9 | 8.15 | 8.5125 | 8.0167 | 8.4667 | 7.1667 | 8.15 | 7.9 | 8.1167 | 8 | 8.1333 | 5.1 | 8.295 | 10.4833 | 8.0667
| Median |  | 8.2407 | 8.1 | 8.05 | 7.3357 | 8 | 7.85 | 8.1 | 8.8 | 8.2 | 8.85 | 8.15 | 8.5152 | 7.95 | 8.45 | 7 | 8.15 | 7.9 | 8.1 | 8 | 8.15 | 5.1 | 8.365 | 10.6 | 8
| Std Dev. |  | 0.1094 | 0.35777 | 0.41312 | 0.51182 | 0 | 0.1472 | 0.58452 | 0.61536 | 0.15055 | 0.21909 | 0.16432 | 0.092574 | 0.16021 | 0.12111 | 0.40825 | 0.10488 | 0.14142 | 0.13292 | 0 | 0.25033 | 0.063246 | 0.22349 | 0.63061 | 0.19664
| Rel.Std.Dev. |  | 1.326% | 4.472% | 5.121% | 7.083% | 0% | 1.891% | 7.028% | 7.019% | 1.844% | 2.462% | 2.016% | 1.088% | 1.998% | 1.43% | 5.696% | 1.287% | 1.79% | 1.638% | 0% | 3.078% | 1.24% | 2.694% | 6.015% | 2.438%
| PDM3 |  | 1.172% | -1.926% | -1.108% | -11.42% | -1.926% | -4.582% | 1.957% | 7.473% | 0.1176% | 9.108% | -0.08669% | 4.357% | -1.721% | 3.795% | -12.14% | -0.08669% | -3.152% | -0.4953% | -1.926% | -0.291% | -37.48% | 1.691% | 28.52% | -1.108%
| Z-Score (Absolute) |  | 0.46 | 0.67 | 0.37 | 4.15 | NA | 1.65 | 0.75 | 2.77 | 0.07 | 3.37 | 0 | 1.63 | 0.6 | 1.42 | NA | 0 | 1.12 | 0.15 | NA | 0.07 | 13.71 | 0.65 | 10.49 | 0.37
Results from laboratories 1.11, 1.22 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 284 (Certified Value 0.4 ppm) | [30 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.29 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | < 1.5 | < 1 | 0.71207 | < 5 | 0.6 | < 5 | 0.35 | 0.3 | 0.4 | 0.4 | 0.7 | 0.6 | 0.4 | < 1 | 2.8 | < 1 | < 1 | 0.3 | 0.3 | < 1 | < 0.1 | 1 | < 1 | [6 empty] | 1
| 1 | 2 | < 1.5 | < 1 | 0.76102 | < 5 | 0.7 | < 5 | 0.33 | 0.5 | 0.6 | 0.4 | 0.2 | 0.5 | 0.7 | < 1 | 3.2 | < 1 | < 1 | 0.4 | 0.3 | < 1 | < 0.1 | 1.1 | < 1 | [6 empty] | 38
| 1 | 3 | < 1.5 | < 1 | 0.87183 | < 5 | 0.8 | < 5 | 0.34 | 0.2 | 0.4 | 0.3 | 0.2 | 0.4 | 0.5 | < 1 | 3.5 | < 1 | < 1 | 0.3 | 0.4 | < 1 | < 0.1 | 1 | < 1 | [6 empty] | 16
| 1 | 4 | < 1.5 | < 1 | 0.78295 | < 5 | 0.7 | < 5 | 0.35 | 0.5 | 0.4 | 0.4 | 0.4 | 0.2 | 0.5 | < 1 | 3.1 | < 1 | < 1 | 0.4 | 0.4 | < 1 | < 0.1 | 1.4 | < 1 | [6 empty] | 0.39741
| 1 | 5 | < 1.5 | < 1 | 0.83451 | < 5 | 0.7 | < 5 | 0.34 | 0.4 | 0.3 | 0.4 | 0.4 | 0.6 | 0.7 | < 1 | 2.8 | < 1 | < 1 | 0.3 | 0.4 | < 1 | < 0.1 | 0.7 | < 1 | [6 empty] | 112
| 1 | 6 | < 1.5 | < 1 | 0.94367 | < 5 | 0.6 | < 5 | 0.35 | 0.3 | 0.5 | 0.3 | 0.3 | 0.5 | 0.3 | < 1 | 3 | < 1 | < 1 | 0.3 | 0.5 | < 1 | < 0.1 | 0.9 | < 1
| Mean |  | 1.05: 0.81768; 1.09: 0.68333; 1.12: 0.34333; 1.13: 0.36667; 1.14: 0.43333; 1.15: 0.36667; 1.16: 0.36667; 1.17: 0.46667; 1.18: 0.51667; 1.20: 3.0667; 1.23: 0.33333; 1.24: 0.38333; 1.31: 1.0167
| Median |  | 1.05: 0.80873; 1.09: 0.7; 1.12: 0.345; 1.13: 0.35; 1.14: 0.4; 1.15: 0.4; 1.16: 0.35; 1.17: 0.5; 1.18: 0.5; 1.20: 3.05; 1.23: 0.3; 1.24: 0.4; 1.31: 1
| Std Dev. |  | 1.05: 0.083259; 1.09: 0.075277; 1.12: 0.008165; 1.13: 0.12111; 1.14: 0.10328; 1.15: 0.05164; 1.16: 0.18619; 1.17: 0.15055; 1.18: 0.16021; 1.20: 0.26583; 1.23: 0.05164; 1.24: 0.075277; 1.31: 0.23166
| Rel.Std.Dev. |  | 1.05: 10.18%; 1.09: 11.02%; 1.12: 2.378%; 1.13: 33.03%; 1.14: 23.83%; 1.15: 14.08%; 1.16: 50.78%; 1.17: 32.26%; 1.18: 31.01%; 1.20: 8.668%; 1.23: 15.49%; 1.24: 19.64%; 1.31: 22.79%
| PDM3 |  | 1.05: 105.8%; 1.09: 71.95%; 1.12: -13.61%; 1.13: -7.735%; 1.14: 9.04%; 1.15: -7.735%; 1.16: -7.735%; 1.17: 17.43%; 1.18: 30.01%; 1.20: 671.7%; 1.23: -16.12%; 1.24: -3.541%; 1.31: 155.8%
| Z-Score (Absolute) |  | 1.26 | 0 | 1.61 | 10.11 | 0.93 | 10.11 | 0.79 | 0.67 | 0.34 | 0.67 | 0.67 | 0.17 | 0.08 | 0 | 12.98 | 0 | 0 | 0.84 | 0.59 | 0 | 2.28 | 2.61 | 0
| Analytical results for Si in OREAS 284 (Indicative Value 0.133 wt.%) | [30 empty] | Indicative
Si | Round | Replicate | 1.03: Lab | [6 empty] | 1
| No | No | 1.03: 1.30 | [6 empty] | wt.%
|  |  | 1.03: AR*OES | [6 empty] | 3
|  |  | 1.03: 0.2g | [6 empty] | 3
| 1 | 1 | 1.03: 0.1666 | [6 empty] | 1
| 1 | 2 | 1.03: 0.1455 | [6 empty] | 7
| 1 | 3 | 1.03: 0.1349 | [6 empty] | 16
| 1 | 4 | 1.03: 0.1115 | [6 empty] | 0.13345
| 1 | 5 | 1.03: 0.1231 | [6 empty] | 13
| 1 | 6 | 1.03: 0.1191
| Mean |  | 1.03: 0.13345
| Median |  | 1.03: 0.129
| Std Dev. |  | 1.03: 0.020204
| Rel.Std.Dev. |  | 1.03: 15.14%
| PDM3 |  | 1.03: -2.22e-14%
Z-Score (Absolute) | bar: 1.03
| Analytical results for Sm in OREAS 284 (Indicative Value 4.37 ppm) | [30 empty] | Indicative
Sm | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab | [6 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.29 | [6 empty] | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*MS | [6 empty] | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.5g | [6 empty] | 2
| 1 | 1 | 1.03: 4.8543; 1.04: 4.265; 1.05: 4.7; 1.07: 5.6468; 1.09: 2.66 | [6 empty] | 1
| 1 | 2 | 1.03: 4.9977; 1.04: 4.12; 1.05: 4.5; 1.07: 5.939; 1.09: 2.5 | [6 empty] | 4
| 1 | 3 | 1.03: 4.8232; 1.04: 4.024; 1.05: 4.5; 1.07: 5.6113; 1.09: 2.6 | [6 empty] | 16
| 1 | 4 | 1.03: 4.7121; 1.04: 3.896; 1.05: 4.2; 1.07: 5.7377; 1.09: 2.6 | [6 empty] | 4.3681
| 1 | 5 | 1.03: 5.0638; 1.04: 3.965; 1.05: 5; 1.07: 6.0736; 1.09: 2.5 | [6 empty] | 10
| 1 | 6 | 1.03: 5.1329; 1.04: 3.517; 1.05: 4.5; 1.07: 6.0142; 1.09: 2.39
| Mean |  | 1.03: 4.9307; 1.04: 3.9645; 1.05: 4.5667; 1.07: 5.8371; 1.09: 2.5417
| Median |  | 1.03: 4.926; 1.04: 3.9945; 1.05: 4.5; 1.07: 5.8384; 1.09: 2.55
| Std Dev. |  | 1.03: 0.16017; 1.04: 0.25415; 1.05: 0.26583; 1.07: 0.19736; 1.09: 0.097245
| Rel.Std.Dev. |  | 1.03: 3.248%; 1.04: 6.411%; 1.05: 5.821%; 1.07: 3.381%; 1.09: 3.826%
| PDM3 |  | 1.03: 12.88%; 1.04: -9.24%; 1.05: 4.545%; 1.07: 33.63%; 1.09: -41.81%
| Z-Score (Absolute) |  | 1.03: 0.41; 1.04: 0.67; 1.05: 0; 1.07: 1.42; 1.09: 2.27
| Analytical results for Sn in OREAS 284 (Certified Value 3.54 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 3.617 | 3.64 | 3.8972 | < 5 | 3.37 | < 10 | 3.7 | 3.7 | 3.7 | 3.6 | 3.3 | 3.4288 | 3.4 | 3.1 | 3.4 | 3.6 | 3.6 | 3.5 | 2.82 | < 0.1 | 2.3 | 3.7 | [7 empty] | 1
| 1 | 2 | 3.6164 | 3.61 | 3.9592 | < 5 | 3.23 | < 10 | 3.6 | 3.7 | 3.5 | 3.6 | 3.5 | 3.6456 | 3.4 | 3.2 | 3.3 | 3.6 | 3.4 | 3.78 | 2.81 | < 0.1 | 2.1 | 3.8 | [7 empty] | 13
| 1 | 3 | 3.5347 | 3.69 | 3.7116 | < 5 | 3.36 | < 10 | 3.7 | 3.3 | 3.6 | 3.5 | 3.3 | 3.4773 | 3.4 | 3.4 | 3.2 | 3.6 | 3.4 | 3.85 | 2.82 | < 0.1 | 2.2 | 3.9 | [7 empty] | 16
| 1 | 4 | 3.6238 | 3.56 | 3.5853 | < 5 | 3.27 | < 10 | 3.7 | 3.8 | 3.6 | 3.7 | 3.3 | 3.7263 | 3.4 | 3.2 | 3.3 | 3.5 | 3.5 | 3.72 | 2.86 | < 0.1 | 2.2 | 3.8 | [7 empty] | 3.5449
| 1 | 5 | 3.6232 | 3.59 | 3.6026 | < 5 | 3.36 | < 10 | 3.6 | 3.7 | 3.8 | 3.7 | 3.4 | 3.5613 | 3.4 | 3.1 | 3.2 | 3.6 | 3.5 | 3.9 | 2.79 | < 0.1 | 2.1 | 3.9 | [7 empty] | 113
| 1 | 6 | 3.6351 | 3.66 | 3.741 | < 5 | 3.21 | < 10 | 3.3 | 3.4 | 3.6 | 3.6 | 3.4 | 3.6189 | 3.4 | 3.1 | 3.4 | 3.5 | 3.5 | 3.71 | 2.8 | < 0.1 | 2.1 | 4
| Mean |  | 3.6084 | 3.625 | 3.7495 |  | 3.3 |  | 3.6 | 3.6 | 3.6333 | 3.6167 | 3.3667 | 3.5764 | 3.4 | 3.1833 | 3.3 | 3.5667 | 3.4833 | 3.7433 | 2.8167 |  | 2.1667 | 3.85
| Median |  | 3.6201 | 3.625 | 3.7263 |  | 3.315 |  | 3.65 | 3.7 | 3.6 | 3.6 | 3.35 | 3.5901 | 3.4 | 3.15 | 3.3 | 3.6 | 3.5 | 3.75 | 2.815 |  | 2.15 | 3.85
| Std Dev. |  | 0.036719 | 0.047645 | 0.15221 |  | 0.072111 |  | 0.15492 | 0.2 | 0.10328 | 0.075277 | 0.08165 | 0.11038 | 0 | 0.1169 | 0.089443 | 0.05164 | 0.075277 | 0.1401 | 0.024221 |  | 0.08165 | 0.10488
| Rel.Std.Dev. |  | 1.018% | 1.314% | 4.06% |  | 2.185% |  | 4.303% | 5.556% | 2.843% | 2.081% | 2.425% | 3.086% | 0% | 3.672% | 2.71% | 1.448% | 2.161% | 3.743% | 0.8599% |  | 3.768% | 2.724%
| PDM3 |  | 1.791% | 2.261% | 5.773% |  | -6.907% |  | 1.556% | 1.556% | 2.496% | 2.026% | -5.027% | 0.8887% | -4.086% | -10.2% | -6.907% | 0.6153% | -1.736% | 5.599% | -20.54% |  | -38.88% | 8.608%
| Z-Score (Absolute) |  | 0.14 | 0.21 | 0.69 | 4.13 | 1.05 | 5.51 | 0.11 | 0.11 | 0.24 | 0.17 | 0.79 | 0.02 | 0.66 | 1.5 | 1.05 | 0.02 | 0.34 | 0.66 | 2.91 | 13.58 | 5.42 | 1.07
| Analytical results for Sr in OREAS 284 (Certified Value 19.3 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 20.84 | 19.72 | 28.4598 | 22.8883 | 31.1 | 7.8101 | < 5 | 18.1 | 21.9 | 24 | 21.7 | 20.9 | 21.2 | 18.2405 | 20 | 17.3 | 18.9 | 17.7 | 19.1 | 19.4 | 18 | 15.67 | 12.84 | 16 | 31.35 | 18.4 | 21 |  |  | 1
| 1 | 2 | 21.7676 | 20.02 | 28.5074 | 22.5981 | 29.1 | 7.4477 | < 5 | 19.2 | 21 | 23.6 | 22.2 | 20.5 | 22.2 | 20.1316 | 20 | 18.3 | 18.7 | 17.8 | 19.8 | 18.5 | 17 | 17.27 | 12.12 | 15 | 31.17 | 17.5 | 21.4 |  |  | 14
| 1 | 3 | 21.422 | 19.11 | 28.5652 | 23.5507 | 29.7 | 8.0186 | < 5 | 19.1 | 20.7 | 21.6 | 21.6 | 20.6 | 21.4 | 20.362 | 20 | 17.5 | 18.8 | 17.9 | 18.8 | 19 | 18 | 17.38 |  | 15 | 30.4 | 16.8 | 21.8 |  |  | 16
| 1 | 4 | 21.1626 | 19.15 | 26.4544 | 22.4843 | 29.9 | 7.0945 | < 5 | 19.3 | 20.7 | 24.3 | 20.6 | 21.3 | 20.9 | 18.9727 | 20 | 18 | 18.7 | 17.5 | 19.5 | 19.2 | 17 | 16.63 | 12.34 | 16 | 26.78 | 18.7 | 21.8 |  |  | 19.313
| 1 | 5 | 22.3742 | 19.28 | 28.975 | 23.4653 | 29.9 | 8.4984 | < 5 | 19.5 | 21.3 | 23 | 22.5 | 21.3 | 21.6 | 21.1843 | 20 | 17.3 | 19 | 17.7 | 19.7 | 19.1 | 17 | 17.29 | 14.46 | 15 | 30.38 | 15.6 | 21.9 |  |  | 114
| 1 | 6 | 21.0539 | 20.36 | 28.8235 | 23.9757 | 28.7 | 8.2127 | < 5 | 18.1 | 21 | 20.1 | 22.2 | 20.8 | 21.3 | 19.563 | 21 | 18.1 | 18.9 | 17.7 | 19.3 | 18.9 | 17 | 16.94 | 12.32 | 15 | 29.85 | 17 | 22.3
| Mean |  | 21.4367 | 19.6067 | 28.2975 | 23.1604 | 29.7333 | 7.847 |  | 18.8833 | 21.1 | 22.7667 | 21.8 | 20.9 | 21.4333 | 19.7424 | 20.1667 | 17.75 | 18.8333 | 17.7167 | 19.3667 | 19.0167 | 17.3333 | 16.8633 | 12.816 | 15.3333 | 29.9883 | 17.3333 | 21.7
| Median |  | 21.2923 | 19.5 | 28.5363 | 23.1768 | 29.8 | 7.9143 |  | 19.15 | 21 | 23.3 | 21.95 | 20.85 | 21.35 | 19.8473 | 20 | 17.75 | 18.85 | 17.7 | 19.4 | 19.05 | 17 | 17.105 | 12.34 | 15 | 30.39 | 17.25 | 21.8
| Std Dev. |  | 0.55968 | 0.51247 | 0.92465 | 0.59285 | 0.82381 | 0.51275 |  | 0.62102 | 0.45166 | 1.6182 | 0.67823 | 0.34059 | 0.44121 | 1.0483 | 0.40825 | 0.43704 | 0.12111 | 0.13292 | 0.37771 | 0.30605 | 0.5164 | 0.64788 | 0.9566 | 0.5164 | 1.6665 | 1.1343 | 0.44721
| Rel.Std.Dev. |  | 2.611% | 2.614% | 3.268% | 2.56% | 2.771% | 6.534% |  | 3.289% | 2.141% | 7.108% | 3.111% | 1.63% | 2.059% | 5.31% | 2.024% | 2.462% | 0.643% | 0.7502% | 1.95% | 1.609% | 2.979% | 3.842% | 7.464% | 3.368% | 5.557% | 6.544% | 2.061%
| PDM3 |  | 11% | 1.52% | 46.52% | 19.92% | 53.95% | -59.37% |  | -2.225% | 9.253% | 17.88% | 12.88% | 8.217% | 10.98% | 2.223% | 4.42% | -8.093% | -2.484% | -8.266% | 0.2778% | -1.534% | -10.25% | -12.68% | -33.64% | -20.61% | 55.28% | -10.25% | 12.36%
| Z-Score (Absolute) |  | 0.59 | 0 | 2.8 | 1.14 | 3.26 | 3.79 | 5.51 | 0.23 | 0.48 | 1.02 | 0.71 | 0.42 | 0.59 | 0.04 | 0.18 | 0.6 | 0.25 | 0.61 | 0.08 | 0.19 | 0.73 | 0.88 | 2.19 | 1.38 | 3.34 | 0.73 | 0.67
| Analytical results for Ta in OREAS 284 (Certified Value < 0.01 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | [11 empty] | 3
| 1 | 1 | < 0.01 | < 0.01 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 1
| 1 | 2 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 25
| 1 | 3 | < 0.01 | < 0.01 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 16
| 1 | 4 | < 0.01 | < 0.01 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 115
| 1 | 6 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05
| Mean |  | 1.09: 0.01; 1.10: 0.01; 1.13: 0.01; 1.14: 0.01
| Median |  | 1.09: 0.01; 1.10: 0.01; 1.13: 0.01; 1.14: 0.01
| Std Dev. |  | 1.09: 0; 1.13: 0; 1.14: 0
Rel.Std.Dev. | bar: 1.09; 1.13 to 1.14
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 2.02 | 335.81 | 0.34 | 0.56 | 0.67 | 0 | 0 | 2.02 | 5.39 | 2.02 | 2.02 | 0.67 | 2.02 | 0.67 | 0.67 | 2.02
| Analytical results for Tb in OREAS 284 (Certified Value 0.57 ppm) | [30 empty] | Certified
Tb | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.13: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.07: 1.10; 1.09: 1.21; 1.10: 1.25; 1.11: 1.29; 1.13: 1.32 |  |  | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*MS; 1.13: AR*OES/MS |  |  | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.07: 0.5g; 1.09: 0.25g; 1.10: 0.25g; 1.11: 0.5g; 1.13: 0.25g |  |  | 3
| 1 | 1 | 1.03: 0.61519; 1.04: 0.526; 1.05: 0.6; 1.07: 0.6421; 1.09: 0.53; 1.10: 0.57; 1.11: 0.38; 1.13: 0.53 |  |  | 1
| 1 | 2 | 1.03: 0.61968; 1.04: 0.524; 1.05: 0.5; 1.07: 0.67; 1.09: 0.55; 1.10: 0.56; 1.11: 0.36; 1.13: 0.53 |  |  | 26
| 1 | 3 | 1.03: 0.5966; 1.04: 0.522; 1.05: 0.6; 1.07: 0.6399; 1.09: 0.55; 1.10: 0.56; 1.11: 0.37; 1.13: 0.54 |  |  | 16
| 1 | 4 | 1.03: 0.59202; 1.04: 0.507; 1.05: 0.5; 1.07: 0.6422; 1.09: 0.54; 1.10: 0.57; 1.11: 0.36; 1.13: 0.56 |  |  | 0.57463
| 1 | 5 | 1.03: 0.61962; 1.04: 0.517; 1.05: 0.6; 1.07: 0.6858; 1.09: 0.53; 1.10: 0.59; 1.11: 0.36; 1.13: 0.53 |  |  | 116
| 1 | 6 | 1.03: 0.61387; 1.04: 0.531; 1.05: 0.6; 1.07: 0.6926; 1.09: 0.53; 1.10: 0.6; 1.11: 0.34; 1.13: 0.56
| Mean |  | 1.03: 0.6095; 1.04: 0.52117; 1.05: 0.56667; 1.07: 0.6621; 1.09: 0.53833; 1.10: 0.575; 1.11: 0.36167; 1.13: 0.54167
| Median |  | 1.03: 0.61453; 1.04: 0.523; 1.05: 0.6; 1.07: 0.6561; 1.09: 0.535; 1.10: 0.57; 1.11: 0.36; 1.13: 0.535
| Std Dev. |  | 1.03: 0.01208; 1.04: 0.0083287; 1.05: 0.05164; 1.07: 0.023846; 1.09: 0.0098319; 1.10: 0.016432; 1.11: 0.013292; 1.13: 0.01472
| Rel.Std.Dev. |  | 1.03: 1.982%; 1.04: 1.598%; 1.05: 9.113%; 1.07: 3.602%; 1.09: 1.826%; 1.10: 2.858%; 1.11: 3.675%; 1.13: 2.717%
| PDM3 |  | 1.03: 6.068%; 1.04: -9.303%; 1.05: -1.385%; 1.07: 15.22%; 1.09: -6.316%; 1.10: 0.0649%; 1.11: -37.06%; 1.13: -5.736%
| Z-Score (Absolute) |  | 1.03: 1.37; 1.04: 0.41; 1.05: NA; 1.07: 2.44; 1.09: 0.07; 1.10: 0.67; 1.11: 3.64; 1.13: 0
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 284 (Certified Value 0.049 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.29 | 1.30 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | [6 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.2 | < 0.1 | 0.04 | 5 | 0.05 | 0.07 | 0.05 | 0.04 | 0.04 | 0.05 | 0.06 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.06 | 0.05 | < 0.05 | < 2 | 0.04 | < 0.2 | < 0.01 | 0.1 | [6 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.04 | 5 | 0.04 | 0.03 | 0.06 | 0.04 | 0.03 | 0.05 | 0.09 | < 0.05 | < 0.1 | < 0.05 | 0.06 | 0.03 | 0.06 | 0.05 | < 2 | 0.05 | < 0.2 | < 0.01 | 0.07 | [6 empty] | 27
| 1 | 3 | < 0.2 | < 0.1 | 0.06 | 5 | 0.04 | 0.06 | 0.04 | 0.03 | 0.01 | 0.05 | 0.08 | < 0.05 | < 0.1 | < 0.05 | 0.07 | 0.05 | 0.04 | < 0.05 | < 2 | 0.05 | < 0.2 | < 0.01 | 0.06 | [6 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.04 | < 5 | 0.06 | 0.06 | 0.05 | 0.05 | 0.06 | 0.04 | 0.02 | < 0.05 | < 0.1 | < 0.05 | 0.07 | 0.05 | 0.06 | < 0.05 | < 2 | 0.04 | < 0.2 | 0.03 | 0.06 | [6 empty] | 0.049128
| 1 | 5 | < 0.2 | < 0.1 | 0.03 | < 5 | 0.06 | 0.06 | 0.01 | 0.03 | 0.03 | 0.06 | 0.06 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.03 | 0.06 | < 0.05 | < 2 | 0.04 | < 0.2 | 0.12 | 0.06 | [6 empty] | 117
| 1 | 6 | < 0.2 | < 0.1 | 0.04 | < 5 | 0.05 | 0.04 | 0.06 | 0.04 | 0.03 | 0.04 | 0.04 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.07 | 0.06 | < 0.05 | < 2 | 0.04 | < 0.2 | < 0.01 | 0.08
| Mean |  |  |  | 0.041667 | 5 | 0.05 | 0.053333 | 0.045 | 0.038333 | 0.033333 | 0.048333 | 0.058333 |  |  |  | 0.066667 | 0.048333 | 0.055 | 0.05 |  | 0.043333 |  | 0.075 | 0.071667
| Median |  |  |  | 0.04 | 5 | 0.05 | 0.06 | 0.05 | 0.04 | 0.03 | 0.05 | 0.06 |  |  |  | 0.07 | 0.05 | 0.06 | 0.05 |  | 0.04 |  | 0.075 | 0.065
| Std Dev. |  |  |  | 0.0098319 | 0 | 0.0089443 | 0.015055 | 0.018708 | 0.0075277 | 0.01633 | 0.0075277 | 0.025626 |  |  |  | 0.0057735 | 0.016021 | 0.0083666 |  |  | 0.005164 |  | 0.06364 | 0.016021
| Rel.Std.Dev. |  |  |  | 23.6% | 0% | 17.89% | 28.23% | 41.57% | 19.64% | 48.99% | 15.57% | 43.93% |  |  |  | 8.66% | 33.15% | 15.21% |  |  | 11.92% |  | 84.85% | 22.35%
| PDM3 |  |  |  | -15.19% | 10077% | 1.775% | 8.559% | -8.403% | -21.97% | -32.15% | -1.618% | 18.74% |  |  |  | 35.7% | -1.618% | 11.95% | 1.775% |  | -11.8% |  | 52.66% | 45.88%
| Z-Score (Absolute) |  | 4.18 | 0.13 | 0.54 | NA | 0.13 | 0.4 | 0.27 | 0.81 | 1.21 | 0 | 0.81 | 1.89 | 0.13 | 1.89 | 0.2 | 0 | 0.54 | 1.55 | 77.01 | 0.4 | 4.18 | 1.62 | 1.89
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 284 (Certified Value 8.52 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.09 | 1.10 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 9.3001 | 8.07 | 6.6277 | 10.5 | 8.5804 | 6.9 | 8.5 | 9.2 | 7.7 | 7.8 | 7.2 | 7.7454 | 7.4 | 6.4 | 11.7 | 6.6 | 5.8 | 9 | 9.22 | 3.76 | 113.2 | 12.4 | 9.5 | [6 empty] | 1
| 1 | 2 | 9.155 | 7.78 | 7.6765 | 9.6 | 10.1424 | 7.5 | 8.1 | 9.7 | 7.5 | 7.3 | 7.4 | 7.7326 | 7.9 | 6.7 | 11.8 | 6.5 | 6.2 | 9.3 | 10.29 | 3.24 | 110.24 | 11.6 | 9.5 | [6 empty] | 28
| 1 | 3 | 9.0549 | 7.5 | 7.4811 | 9.8 | 9.691 | 8.4 | 8.5 | 9 | 7.5 | 7.6 | 7.1 | 7.9088 | 8 | 6.7 | 11.6 | 6.7 | 5.1 | 9.1 | 10.55 | 3.42 | 107.6 | 11 | 9.4 | [6 empty] | 16
| 1 | 4 | 9.2323 | 7.22 | 6.6582 | 9.8 | 10.174 | 8.1 | 8.4 | 10 | 7 | 7.7 | 7 | 8.371 | 8.1 | 6.5 | 11.8 | 6.4 | 5.5 | 9.5 | 9.86 | 3.32 | 101.04 | 11.9 | 9.7 | [6 empty] | 8.5151
| 1 | 5 | 9.0805 | 7.28 | 7.1707 | 10.4 | 10.1234 | 7.9 | 8.6 | 9.5 | 7.7 | 8 | 7.5 | 7.7178 | 8.2 | 6.3 | 11.7 | 6.6 | 6 | 9.7 | 10.13 | 3.19 | 107.92 | 10.1 | 9.4 | [6 empty] | 118
| 1 | 6 | 9.394 | 6.44 | 7.9894 | 10.3 | 10.0366 | 7.4 | 6.9 | 7.7 | 7.6 | 7.8 | 7.3 | 8.3596 | 8.4 | 6.6 | 12 | 6.6 | 5.5 | 9.9 | 10.28 | 3.02 | 105.76 | 10.9 | 9.7
| Mean |  | 9.2028 | 7.3817 | 7.2673 | 10.0667 | 9.7913 | 7.7 | 8.1667 | 9.1833 | 7.5 | 7.7 | 7.25 | 7.9725 | 8 | 6.5333 | 11.7667 | 6.5667 | 5.6833 | 9.4167 | 10.055 | 3.325 | 107.6267 | 11.3167 | 9.5333
| Median |  | 9.1936 | 7.39 | 7.3259 | 10.05 | 10.08 | 7.7 | 8.45 | 9.35 | 7.55 | 7.75 | 7.25 | 7.8271 | 8.05 | 6.55 | 11.75 | 6.6 | 5.65 | 9.4 | 10.205 | 3.28 | 107.76 | 11.3 | 9.5
| Std Dev. |  | 0.13114 | 0.5603 | 0.5521 | 0.37771 | 0.61912 | 0.54037 | 0.64395 | 0.8085 | 0.26077 | 0.23664 | 0.18708 | 0.31199 | 0.34059 | 0.1633 | 0.13663 | 0.10328 | 0.39707 | 0.34881 | 0.46728 | 0.25169 | 4.1194 | 0.81833 | 0.13663
| Rel.Std.Dev. |  | 1.425% | 7.59% | 7.597% | 3.752% | 6.323% | 7.018% | 7.885% | 8.804% | 3.477% | 3.073% | 2.58% | 3.913% | 4.257% | 2.499% | 1.161% | 1.573% | 6.987% | 3.704% | 4.647% | 7.57% | 3.827% | 7.231% | 1.433%
| PDM3 |  | 8.076% | -13.31% | -14.65% | 18.22% | 14.99% | -9.572% | -4.092% | 7.848% | -11.92% | -9.572% | -14.86% | -6.371% | -6.049% | -23.27% | 38.19% | -22.88% | -33.26% | 10.59% | 18.08% | -60.95% | 1164% | 32.9% | 11.96%
| Z-Score (Absolute) |  | 0.62 | 0.31 | 0.37 | 1.06 | 0.92 | 0.15 | 0.09 | 0.61 | 0.25 | 0.15 | 0.38 | 0.01 | 0.01 | 0.74 | 1.93 | 0.73 | 1.18 | 0.73 | 1.06 | 2.38 | NA | 1.7 | 0.79
| Analytical results for Ti in OREAS 284 (Certified Value 0.273 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.26127 | 0.2681 | 0.29941 | 0.30715 | 0.23 | 0.3015 | 0.25 | 0.246 | 0.291 | 0.297 | 0.278 | 0.278 | 0.283 | 0.26412 | 0.231 | 0.29 | 0.31 | 0.27 | 0.246 | 0.27 | 0.2538 | 0.2568 | 0.343 | 0.289 | 0.25632 | 0.388 | 0.29 |  |  | 1
| 1 | 2 | 0.25416 | 0.2645 | 0.30285 | 0.2924 | 0.23 | 0.3175 | 0.25 | 0.259 | 0.298 | 0.3 | 0.282 | 0.275 | 0.282 | 0.26323 | 0.231 | 0.3 | 0.3 | 0.28 | 0.251 | 0.25 | 0.2495 | 0.272 | 0.358 | 0.277 | 0.2508 | 0.373 | 0.3 |  |  | 29
| 1 | 3 | 0.25154 | 0.2611 | 0.29499 | 0.30029 | 0.22 | 0.2806 | 0.25 | 0.257 | 0.298 | 0.29 | 0.275 | 0.28 | 0.281 | 0.28059 | 0.234 | 0.29 | 0.302 | 0.27 | 0.244 | 0.26 | 0.2539 | 0.282 | 0.365 | 0.28 | 0.24536 | 0.356 | 0.3 |  |  | 16
| 1 | 4 | 0.2515 | 0.2653 | 0.291 | 0.30702 | 0.22 | 0.279 | 0.25 | 0.258 | 0.294 | 0.291 | 0.273 | 0.278 | 0.284 | 0.2773 | 0.236 | 0.3 | 0.308 | 0.27 | 0.257 | 0.27 | 0.2439 | 0.2722 | 0.364 | 0.2866 | 0.23368 | 0.384 | 0.3 |  |  | 0.27265
| 1 | 5 | 0.25322 | 0.2593 | 0.30495 | 0.30437 | 0.22 | 0.2821 | 0.25 | 0.259 | 0.303 | 0.294 | 0.279 | 0.283 | 0.289 | 0.27715 | 0.238 | 0.3 | 0.3 | 0.28 | 0.255 | 0.26 | 0.2506 | 0.287 | 0.377 | 0.2842 | 0.25544 | 0.324 | 0.3 |  |  | 119
| 1 | 6 | 0.25979 | 0.2706 | 0.3041 | 0.30186 | 0.22 | 0.2881 | 0.25 | 0.246 | 0.279 | 0.276 | 0.278 | 0.277 | 0.295 | 0.27795 | 0.236 | 0.29 | 0.302 | 0.27 | 0.256 | 0.26 | 0.2392 | 0.2751 | 0.358 | 0.2806 | 0.25304 | 0.346 | 0.31
| Mean |  | 0.25525 | 0.26482 | 0.29955 | 0.30218 | 0.22333 | 0.29147 | 0.25 | 0.25417 | 0.29383 | 0.29133 | 0.2775 | 0.2785 | 0.28567 | 0.27339 | 0.23433 | 0.295 | 0.30367 | 0.27333 | 0.2515 | 0.26167 | 0.24848 | 0.27418 | 0.36083 | 0.2829 | 0.24911 | 0.36183 | 0.3
| Median |  | 0.25369 | 0.2649 | 0.30113 | 0.30312 | 0.22 | 0.2851 | 0.25 | 0.2575 | 0.296 | 0.2925 | 0.278 | 0.278 | 0.2835 | 0.27723 | 0.235 | 0.295 | 0.302 | 0.27 | 0.253 | 0.26 | 0.25005 | 0.27365 | 0.361 | 0.2824 | 0.25192 | 0.3645 | 0.3
| Std Dev. |  | 0.004242 | 0.0042145 | 0.0055634 | 0.0055177 | 0.005164 | 0.015172 | 0 | 0.0063692 | 0.0083287 | 0.0083825 | 0.0031464 | 0.0027386 | 0.0053541 | 0.0076305 | 0.0028752 | 0.0054772 | 0.004274 | 0.005164 | 0.0054681 | 0.0075277 | 0.0058362 | 0.010364 | 0.011161 | 0.0044967 | 0.0085147 | 0.024564 | 0.0063246
| Rel.Std.Dev. |  | 1.662% | 1.591% | 1.857% | 1.826% | 2.312% | 5.205% | 0% | 2.506% | 2.834% | 2.877% | 1.134% | 0.9833% | 1.874% | 2.791% | 1.227% | 1.857% | 1.407% | 1.889% | 2.174% | 2.877% | 2.349% | 3.78% | 3.093% | 1.589% | 3.418% | 6.789% | 2.108%
| PDM3 |  | -6.384% | -2.874% | 9.865% | 10.83% | -18.09% | 6.9% | -8.309% | -6.78% | 7.768% | 6.851% | 1.778% | 2.144% | 4.773% | 0.2712% | -14.05% | 8.196% | 11.37% | 0.2493% | -7.758% | -4.03% | -8.865% | 0.5611% | 32.34% | 3.758% | -8.636% | 32.71% | 10.03%
| Z-Score (Absolute) |  | 0.68 | 0.39 | 0.67 | 0.75 | 1.66 | 0.43 | 0.84 | 0.71 | 0.5 | 0.42 | 0 | 0.03 | 0.25 | 0.13 | 1.32 | 0.54 | 0.8 | 0.13 | 0.8 | 0.48 | 0.89 | 0.1 | 2.55 | 0.16 | 0.87 | 2.58 | 0.69
| Analytical results for Tl in OREAS 284 (Certified Value 0.68 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 0.7022 | 0.7 | 0.64392 | 0.7 | < 10 | 0.441 | 0.57 | 0.68 | 0.69 | 0.7 | 0.7 | 0.61 | 0.63224 | 0.71 | 0.66 | 0.67 | 0.68 | 0.63 | 0.68 | < 5 | 0.68 | < 5 | 0.73 | 0.61 | [5 empty] | 1
| 1 | 2 | 0.66317 | 0.72 | 0.71126 | 0.68 | < 10 | 0.426 | 0.59 | 0.69 | 0.7 | 0.68 | 0.67 | 0.62 | 0.67403 | 0.69 | 0.68 | 0.64 | 0.68 | 0.63 | 0.65 | < 5 | 0.71 | < 5 | 0.7 | 0.63 | [5 empty] | 30
| 1 | 3 | 0.66461 | 0.72 | 0.69126 | 0.65 | < 10 | 0.433 | 0.65 | 0.72 | 0.6 | 0.71 | 0.67 | 0.61 | 0.62876 | 0.7 | 0.68 | 0.66 | 0.7 | 0.62 | 0.67 | < 5 | 0.72 | < 5 | 0.71 | 0.64 | [5 empty] | 16
| 1 | 4 | 0.6808 | 0.71 | 0.65615 | 0.67 | < 10 | 0.415 | 0.62 | 0.71 | 0.66 | 0.68 | 0.67 | 0.58 | 0.66001 | 0.69 | 0.68 | 0.68 | 0.7 | 0.69 | 0.69 | < 5 | 0.69 | < 5 | 0.73 | 0.65 | [5 empty] | 0.67857
| 1 | 5 | 0.67789 | 0.7 | 0.70555 | 0.68 | < 10 | 0.435 | 0.61 | 0.72 | 0.68 | 0.67 | 0.7 | 0.63 | 0.66826 | 0.67 | 0.65 | 0.68 | 0.69 | 0.68 | 0.68 | < 5 | 0.76 | < 5 | 0.66 | 0.62 | [5 empty] | 120
| 1 | 6 | 0.66795 | 0.71 | 0.68929 | 0.68 | < 10 | 0.436 | 0.59 | 0.61 | 0.62 | 0.67 | 0.68 | 0.58 | 0.64387 | 0.69 | 0.65 | 0.68 | 0.69 | 0.61 | 0.68 | < 5 | 0.69 | < 5 | 0.7 | 0.66
| Mean |  | 0.6761 | 0.71 | 0.6829 | 0.67667 |  | 0.431 | 0.605 | 0.68833 | 0.65833 | 0.685 | 0.68167 | 0.605 | 0.6512 | 0.69167 | 0.66667 | 0.66833 | 0.69 | 0.64333 | 0.675 |  | 0.70833 |  | 0.705 | 0.635
| Median |  | 0.67292 | 0.71 | 0.69027 | 0.68 |  | 0.434 | 0.6 | 0.7 | 0.67 | 0.68 | 0.675 | 0.61 | 0.65194 | 0.69 | 0.67 | 0.675 | 0.69 | 0.63 | 0.68 |  | 0.7 |  | 0.705 | 0.635
| Std Dev. |  | 0.014643 | 0.0089443 | 0.027065 | 0.01633 |  | 0.0092304 | 0.028107 | 0.041673 | 0.040208 | 0.016432 | 0.01472 | 0.020736 | 0.01901 | 0.013292 | 0.015055 | 0.016021 | 0.0089443 | 0.033267 | 0.013784 |  | 0.029269 |  | 0.025884 | 0.018708
| Rel.Std.Dev. |  | 2.166% | 1.26% | 3.963% | 2.413% |  | 2.142% | 4.646% | 6.054% | 6.108% | 2.399% | 2.159% | 3.428% | 2.919% | 1.922% | 2.258% | 2.397% | 1.296% | 5.171% | 2.042% |  | 4.132% |  | 3.672% | 2.946%
| PDM3 |  | -0.3634% | 4.632% | 0.6391% | -0.28% |  | -36.48% | -10.84% | 1.439% | -2.982% | 0.9481% | 0.4568% | -10.84% | -4.034% | 1.931% | -1.754% | -1.508% | 1.685% | -5.192% | -0.5256% |  | 4.387% |  | 3.895% | -6.42%
| Z-Score (Absolute) |  | 0.09 | 0.89 | 0.11 | 0.07 | 124.87 | 7.17 | 2.14 | 0.26 | 0.6 | 0.17 | 0.07 | 2.14 | 0.81 | 0.36 | 0.36 | 0.31 | 0.31 | 1.04 | 0.12 | 52.62 | 0.84 | 52.62 | 0.75 | 1.28
| Analytical results for Tm in OREAS 284 (Indicative Value 0.12 ppm) | [30 empty] | Indicative
Tm | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.09: Lab | [5 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.09: 1.10 | [5 empty] | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.09: AR*MS | [5 empty] | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.05: 0.5g; 1.09: 0.5g | [5 empty] | 2
| 1 | 1 | 1.03: 0.11953; 1.04: 0.113; 1.05: 0.1; 1.09: 0.1266 | [5 empty] | 1
| 1 | 2 | 1.03: 0.12202; 1.04: 0.116; 1.05: 0.1; 1.09: 0.1315 | [5 empty] | 5
| 1 | 3 | 1.03: 0.11581; 1.04: 0.115; 1.05: 0.1; 1.09: 0.1259 | [5 empty] | 16
| 1 | 4 | 1.03: 0.11705; 1.04: 0.113; 1.05: 0.1; 1.09: 0.1629 | [5 empty] | 0.11673
| 1 | 5 | 1.03: 0.12228; 1.04: 0.116; 1.05: 0.1; 1.09: 0.1358 | [5 empty] | 11
| 1 | 6 | 1.03: 0.11954; 1.04: 0.121; 1.05: 0.1; 1.09: 0.1397
| Mean |  | 1.03: 0.11937; 1.04: 0.11567; 1.05: 0.1; 1.09: 0.13707
| Median |  | 1.03: 0.11953; 1.04: 0.1155; 1.05: 0.1; 1.09: 0.13365
| Std Dev. |  | 1.03: 0.0025896; 1.04: 0.0029439; 1.05: 1.5202e-17; 1.09: 0.01372
| Rel.Std.Dev. |  | 1.03: 2.169%; 1.04: 2.545%; 1.05: 1.52e-14%; 1.09: 10.01%
| PDM3 |  | 1.03: 2.259%; 1.04: -0.9148%; 1.05: -14.34%; 1.09: 17.42%
| Z-Score (Absolute) |  | 1.03: 0.13; 1.04: 0.13; 1.05: 1.22; 1.09: 1.36
| Analytical results for U in OREAS 284 (Certified Value 3.47 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 3.7274 | 3.7 | 3.7 | 3.2607 | 2.84 | 2.81 | 3.76 | 3.76 | 3.65 | 3.87 | 3.05 | 3.2454 | 3.4 | 3.36 | 3.73 | 2.99 | 3.2 | 3.39 | 3.43 | 3.18 | < 5 | 3.72 | 3.52 | [6 empty] | 1
| 1 | 2 | 3.1073 | 3.54 | 3.5 | 3.3428 | 2.97 | 2.86 | 3.41 | 3.49 | 3.6 | 3.25 | 3.13 | 3.6399 | 3.7 | 3.73 | 3.74 | 3.12 | 3.3 | 3.14 | 3.38 | 2.86 | < 5 | 4.68 | 3.3 | [6 empty] | 32
| 1 | 3 | 3.7439 | 3.58 | 3.7 | 3.1613 | 3.62 | 3.51 | 3.52 | 3.35 | 3.13 | 3.56 | 2.99 | 3.2573 | 3.6 | 3.52 | 3.91 | 3.55 | 3.3 | 3.38 | 4.05 | 3.18 | < 5 | 4.36 | 3.46 | [6 empty] | 16
| 1 | 4 | 3.9877 | 3.66 | 3.5 | 3.5461 | 2.87 | 3.5 | 3.26 | 3.67 | 3.39 | 3.53 | 3.44 | 3.3916 | 3.4 | 3.39 | 3.75 | 3.48 | 3.4 | 3.62 | 3.6 | 3.36 | < 5 | 4.75 | 4.25 | [6 empty] | 3.4653
| 1 | 5 | 3.5256 | 3.7 | 4.1 | 3.233 | 2.86 | 3.76 | 4.01 | 3.88 | 3.83 | 3.47 | 3.68 | 3.7173 | 3.1 | 3.25 | 3.78 | 3.32 | 3.4 | 3.31 | 4.02 | 2.85 | < 5 | 3.64 | 3.26 | [6 empty] | 121
| 1 | 6 | 3.2734 | 3.59 | 4 | 3.3865 | 2.67 | 3.3 | 2.84 | 3.38 | 4.01 | 3.51 | 2.98 | 3.7466 | 3.9 | 3.44 | 3.76 | 3.7 | 3.1 | 3.23 | 3.44 | 2.58 | < 5 | 3.7 | 3.51
| Mean |  | 3.5609 | 3.6283 | 3.75 | 3.3217 | 2.9717 | 3.29 | 3.4667 | 3.5883 | 3.6017 | 3.5317 | 3.2117 | 3.4997 | 3.5167 | 3.4483 | 3.7783 | 3.36 | 3.2833 | 3.345 | 3.6533 | 3.0017 |  | 4.1417 | 3.55
| Median |  | 3.6265 | 3.625 | 3.7 | 3.3018 | 2.865 | 3.4 | 3.465 | 3.58 | 3.625 | 3.52 | 3.09 | 3.5157 | 3.5 | 3.415 | 3.755 | 3.4 | 3.3 | 3.345 | 3.52 | 3.02 |  | 4.04 | 3.485
| Std Dev. |  | 0.3265 | 0.067651 | 0.251 | 0.13588 | 0.33211 | 0.38179 | 0.40564 | 0.21498 | 0.31256 | 0.19944 | 0.28534 | 0.22939 | 0.27869 | 0.16437 | 0.066758 | 0.26915 | 0.1169 | 0.16453 | 0.30487 | 0.28736 |  | 0.51616 | 0.35978
| Rel.Std.Dev. |  | 9.169% | 1.865% | 6.693% | 4.091% | 11.18% | 11.6% | 11.7% | 5.991% | 8.678% | 5.647% | 8.884% | 6.555% | 7.925% | 4.767% | 1.767% | 8.01% | 3.561% | 4.919% | 8.345% | 9.573% |  | 12.46% | 10.13%
| PDM3 |  | 2.759% | 4.706% | 8.217% | -4.142% | -14.24% | -5.058% | 0.04046% | 3.551% | 3.936% | 1.916% | -7.318% | 0.9932% | 1.483% | -0.4886% | 9.034% | -3.038% | -5.25% | -3.471% | 5.427% | -13.38% |  | 19.52% | 2.445%
| Z-Score (Absolute) |  | 0.27 | 0.56 | 1.1 | 0.78 | 2.32 | 0.92 | 0.14 | 0.39 | 0.45 | 0.14 | 1.26 | 0 | 0.07 | 0.23 | 1.22 | 0.61 | 0.95 | 0.68 | 0.67 | 2.19 | 4.39 | 2.82 | 0.22
| Analytical results for V in OREAS 284 (Certified Value 61 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  | 0
|  |  | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 60.351 | 61 | 65.4976 | 63 | 59.2541 | 61 | 59 | 62 | 62 | 60 | 63 | 59 | 57.6673 | 65 | 58 | 59 | 53.8 | 58 | 62 | 59 | 60 | 65.34 | 54.3 | 62.35 | 70.7 | 58 |  |  |  | 1
| 1 | 2 | 58.2482 | 61 | 63.8523 | 60 | 63.0793 | 62 | 61 | 63 | 62 | 60 | 62 | 59 | 59.0976 | 64 | 58 | 60 | 54.4 | 59 | 59 | 58 | 63 | 66.09 | 52.5 | 61.19 | 69 | 60 |  |  |  | 16
| 1 | 3 | 58.0448 | 62 | 64.5829 | 59 | 55.7329 | 61 | 61 | 63 | 60 | 59 | 62 | 59 | 58.2301 | 64 | 56 | 61 | 53.6 | 58 | 60 | 59 | 65 | 66.47 | 52.6 | 60.07 | 67 | 61 |  |  |  | 16
| 1 | 4 | 59.3933 | 61 | 63.3243 | 59 | 56.0185 | 61 | 62 | 62 | 60 | 59 | 63 | 59 | 57.7689 | 63 | 58 | 61 | 53.1 | 63 | 61 | 57 | 63 | 65.59 | 51.9 | 57.95 | 71 | 60 |  |  |  | 60.7629
| 1 | 5 | 58.8465 | 61 | 63.4572 | 62 | 57.4829 | 61 | 62 | 63 | 61 | 60 | 63 | 60 | 58.2566 | 63 | 57 | 60 | 54.4 | 63 | 62 | 58 | 66 | 67.06 | 52.8 | 62.46 | 62 | 61 |  |  |  | 122
| 1 | 6 | 59.6847 | 62 | 63.726 | 57 | 56.4104 | 61 | 59 | 59 | 60 | 60 | 62 | 62 | 58.9659 | 63 | 57 | 60 | 54 | 63 | 61 | 56 | 63 | 65.54 | 51.1 | 61.5 | 65.1 | 62
| Mean |  | 59.0947 | 61.3333 | 64.0734 | 60 | 57.9963 | 61.1667 | 60.6667 | 62 | 60.8333 | 59.6667 | 62.5 | 59.6667 | 58.3311 | 63.6667 | 57.3333 | 60.1667 | 53.8833 | 60.6667 | 60.8333 | 57.8333 | 63.3333 | 66.015 | 52.5333 | 60.92 | 67.4667 | 60.3333
| Median |  | 59.1199 | 61 | 63.7891 | 59.5 | 56.9466 | 61 | 61 | 62.5 | 60.5 | 60 | 62.5 | 59 | 58.2434 | 63.5 | 57.5 | 60 | 53.9 | 61 | 61 | 58 | 63 | 65.84 | 52.55 | 61.345 | 68 | 60.5
| Std Dev. |  | 0.88272 | 0.5164 | 0.82443 | 2.1909 | 2.8012 | 0.40825 | 1.3663 | 1.5492 | 0.98319 | 0.5164 | 0.54772 | 1.2111 | 0.5938 | 0.8165 | 0.8165 | 0.75277 | 0.49967 | 2.582 | 1.169 | 1.169 | 2.0656 | 0.65826 | 1.0633 | 1.6952 | 3.4927 | 1.3663
| Rel.Std.Dev. |  | 1.494% | 0.842% | 1.287% | 3.651% | 4.83% | 0.6674% | 2.252% | 2.499% | 1.616% | 0.8655% | 0.8764% | 2.03% | 1.018% | 1.282% | 1.424% | 1.251% | 0.9273% | 4.256% | 1.922% | 2.021% | 3.261% | 0.9971% | 2.024% | 2.783% | 5.177% | 2.265%
| PDM3 |  | -2.745% | 0.9387% | 5.448% | -1.256% | -4.553% | 0.6645% | -0.1584% | 2.036% | 0.1159% | -1.804% | 2.859% | -1.804% | -4.002% | 4.779% | -5.644% | -0.9813% | -11.32% | -0.1584% | 0.1159% | -4.821% | 4.23% | 8.644% | -13.54% | 0.2585% | 11.03% | -0.707%
| Z-Score (Absolute) |  | 0.73 | 0.31 | 1.58 | 0.31 | 1.24 | 0.23 | 0 | 0.62 | 0.08 | 0.46 | 0.85 | 0.46 | 1.08 | 1.39 | 1.55 | 0.23 | 3.15 | 0 | 0.08 | 1.32 | 1.24 | 2.48 | 3.78 | 0.12 | 3.16 | 0.15
| Analytical results for W in OREAS 284 (Certified Value 10.1 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 1
| 1 | 1 | 10.1469 | 9.39 | 10.7 | 12.0002 | 13 | 7.53 | 10.5 | 10.95 | 10.8 | 9.64 | 10.95 | 10.2499 | 9.2 | 9.6 | 11 | 11.8 | 7.5 | 11.6 | 10 | 10.14 | 9.4 | 9.952 | 9.13 | 6.3 | [5 empty] | 1
| 1 | 2 | 10.0769 | 9.14 | 10.1 | 13.1753 | 12 | 7.6 | 10.45 | 12.1 | 10.85 | 9.18 | 10.65 | 10.4432 | 9.4 | 9.8 | 11.1 | 11.6 | 8.1 | 11.8 | 10 | 10.96 | 8.8 | 10.584 | 9.28 | 6.9 | [5 empty] | 34
| 1 | 3 | 10.2105 | 9.1 | 10.9 | 12.4307 | 13 | 8.73 | 11.85 | 9.77 | 10.85 | 9.5 | 10.3 | 11.0196 | 9.8 | 9.6 | 11 | 11.8 | 7.2 | 11.1 | 10 | 11.08 | 9.1 | 10.456 | 9.36 | 7.7 | [5 empty] | 16
| 1 | 4 | 10.3449 | 8.86 | 10 | 10.8201 | 12 | 7.98 | 10.9 | 11.75 | 9.77 | 9.58 | 9.86 | 11.7337 | 9.8 | 9.8 | 10.9 | 11.8 | 7.5 | 12 | 10 | 10.55 | 9.2 | 8.328 | 9.21 | 8.2 | [5 empty] | 10.1375
| 1 | 5 | 10.3476 | 8.76 | 11 | 12.3974 | 12 | 8.18 | 10.9 | 11.5 | 11.1 | 9.64 | 10.7 | 11.8581 | 9.9 | 9.3 | 11.1 | 12.5 | 8 | 12.5 | 10 | 11.03 | 9.3 | 10.208 | 9.21 | 7.3 | [5 empty] | 123
| 1 | 6 | 10.2842 | 9.4 | 11 | 11.9823 | 12 | 8.62 | 9.37 | 10.5 | 11.05 | 9.35 | 10.75 | 11.1099 | 9.8 | 9.6 | 11.1 | 11.6 | 7.5 | 12.7 | 9 | 10.9 | 8.7 | 10.392 | 9.51 | 9.4
| Mean |  | 10.2352 | 9.1083 | 10.6167 | 12.1343 | 12.3333 | 8.1067 | 10.6617 | 11.095 | 10.7367 | 9.4817 | 10.535 | 11.0691 | 9.65 | 9.6167 | 11.0333 | 11.85 | 7.6333 | 11.95 | 9.8333 | 10.7767 | 9.0833 | 9.9867 | 9.2833 | 7.6333
| Median |  | 10.2473 | 9.12 | 10.8 | 12.1988 | 12 | 8.08 | 10.7 | 11.225 | 10.85 | 9.54 | 10.675 | 11.0648 | 9.8 | 9.6 | 11.05 | 11.8 | 7.5 | 11.9 | 10 | 10.93 | 9.15 | 10.3 | 9.245 | 7.5
| Std Dev. |  | 0.11001 | 0.264 | 0.4535 | 0.77575 | 0.5164 | 0.50254 | 0.80809 | 0.86415 | 0.48874 | 0.18357 | 0.39236 | 0.65287 | 0.28107 | 0.18348 | 0.08165 | 0.33317 | 0.34448 | 0.58907 | 0.40825 | 0.36379 | 0.27869 | 0.84175 | 0.1353 | 1.0838
| Rel.Std.Dev. |  | 1.075% | 2.898% | 4.272% | 6.393% | 4.187% | 6.199% | 7.579% | 7.789% | 4.552% | 1.936% | 3.724% | 5.898% | 2.913% | 1.908% | 0.74% | 2.812% | 4.513% | 4.929% | 4.152% | 3.376% | 3.068% | 8.429% | 1.457% | 14.2%
| PDM3 |  | 0.9638% | -10.15% | 4.727% | 19.7% | 21.66% | -20.03% | 5.171% | 9.445% | 5.911% | -6.469% | 3.921% | 9.19% | -4.809% | -5.137% | 8.837% | 16.89% | -24.7% | 17.88% | -3% | 6.305% | -10.4% | -1.488% | -8.426% | -24.7%
| Z-Score (Absolute) |  | 0.13 | 1.15 | 0.21 | 1.57 | NA | 2.04 | 0.25 | 0.64 | 0.32 | 0.81 | 0.13 | 0.61 | 0.66 | 0.69 | 0.58 | 1.31 | 2.47 | 1.4 | NA | 0.35 | 1.17 | 0.36 | 0.99 | 2.47
Results from laboratories 1.11 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 284 (Certified Value 10.8 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [8 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [8 empty] | 2
| 1 | 1 | 11.2166 | 11.13 | 11 | 11.362 | 11 | 10.9 | 12 | 10.55 | 11.25 | 10.55 | 10.8894 | 10.3 | 9.42 | 12.8 | 10.1 | 10.75 | 10.73 | 6.84 | 14.994 | 12.4 | 10.6 | [8 empty] | 1
| 1 | 2 | 11.021 | 11.06 | 10.5 | 11.9225 | 11 | 10.6 | 11.6 | 10.4 | 11.15 | 10.9 | 10.3637 | 10.5 | 9.85 | 12.9 | 10.1 | 10.58 | 11.5 | 6.47 | 14.355 | 12 | 10.92 | [8 empty] | 35
| 1 | 3 | 10.8683 | 11.2 | 10.8 | 10.588 | 11 | 11 | 10.4 | 10.25 | 11 | 10.15 | 10.0483 | 10.5 | 9.85 | 12.8 | 10.1 | 10.77 | 11.76 | 6.64 | 13.986 | 12 | 11.05 | [8 empty] | 16
| 1 | 4 | 10.8597 | 10.81 | 10.6 | 10.5199 | 10 | 10.95 | 12.15 | 10.05 | 11.5 | 9.91 | 11.0642 | 10.4 | 9.82 | 12.9 | 9.99 | 10.69 | 11.26 | 6.64 | 13.743 | 12.2 | 11.11 | [8 empty] | 10.8395
| 1 | 5 | 11.1062 | 10.82 | 11 | 10.6729 | 11 | 10.85 | 11.95 | 10.85 | 11.5 | 10.25 | 10.2686 | 10.5 | 9.8 | 12.7 | 10.1 | 11.11 | 11.75 | 6.58 | 14.481 | 11.4 | 11.03 | [8 empty] | 124
| 1 | 6 | 11.2217 | 11.14 | 10.3 | 10.7614 | 10 | 10.9 | 11.1 | 10.5 | 11.15 | 10.4 | 11.122 | 10.6 | 9.94 | 12.8 | 10.1 | 11.26 | 11.49 | 6.34 | 14.382 | 11.7 | 11.36
| Mean |  | 11.0489 | 11.0267 | 10.7 | 10.9711 | 10.6667 | 10.8667 | 11.5333 | 10.4333 | 11.2583 | 10.36 | 10.626 | 10.4667 | 9.78 | 12.8167 | 10.0817 | 10.86 | 11.415 | 6.585 | 14.3235 | 11.95 | 11.0117
| Median |  | 11.0636 | 11.095 | 10.7 | 10.7172 | 11 | 10.9 | 11.775 | 10.45 | 11.2 | 10.325 | 10.6265 | 10.5 | 9.835 | 12.8 | 10.1 | 10.76 | 11.495 | 6.61 | 14.3685 | 12 | 11.04
| Std Dev. |  | 0.16152 | 0.1699 | 0.28284 | 0.5552 | 0.5164 | 0.14024 | 0.67057 | 0.27325 | 0.20351 | 0.34293 | 0.45554 | 0.10328 | 0.18276 | 0.075277 | 0.044907 | 0.26458 | 0.38402 | 0.16991 | 0.4312 | 0.35637 | 0.24927
| Rel.Std.Dev. |  | 1.462% | 1.541% | 2.643% | 5.061% | 4.841% | 1.291% | 5.814% | 2.619% | 1.808% | 3.31% | 4.287% | 0.9867% | 1.869% | 0.5873% | 0.4454% | 2.436% | 3.364% | 2.58% | 3.01% | 2.982% | 2.264%
| PDM3 |  | 1.932% | 1.727% | -1.287% | 1.214% | -1.594% | 0.2508% | 6.401% | -3.747% | 3.864% | -4.423% | -1.969% | -3.439% | -9.774% | 18.24% | -6.991% | 0.1893% | 5.31% | -39.25% | 32.14% | 10.25% | 1.589%
| Z-Score (Absolute) |  | 0.19 | 0.15 | 0.31 | 0.08 | NA | 0.08 | 0.88 | 0.7 | 0.49 | 0.8 | 0.42 | 0.65 | 1.64 | 2.73 | 1.2 | 0.08 | 0.71 | 6.23 | 4.9 | 1.48 | 0.13
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 284 (Certified Value 0.71 ppm) | [30 empty] | Certified
Yb | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab | [8 empty] | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.09: 1.09; 1.10: 1.10; 1.11: 1.21; 1.14: 1.25; 1.15: 1.29; 1.16: 1.32 | [8 empty] | ppm
|  |  | 1.03: AR*MS; 1.04: AR*MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*OES/MS; 1.14: AR*MS; 1.15: AR*MS; 1.16: AR*OES/MS | [8 empty] | 2
|  |  | 1.03: 0.2g; 1.04: 01g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.25g; 1.14: 0.25g; 1.15: 0.5g; 1.16: 0.25g | [8 empty] | 3
| 1 | 1 | 1.03: 0.75365; 1.04: 0.697; 1.09: 0.7; 1.10: 0.8136; 1.11: 0.7; 1.14: 0.7; 1.15: 0.65; 1.16: 0.7 | [8 empty] | 1
| 1 | 2 | 1.03: 0.74174; 1.04: 0.714; 1.09: 0.7; 1.10: 0.8266; 1.11: 0.7; 1.14: 0.7; 1.15: 0.61; 1.16: 0.7 | [8 empty] | 36
| 1 | 3 | 1.03: 0.72143; 1.04: 0.691; 1.09: 0.7; 1.10: 0.8384; 1.11: 0.7; 1.14: 0.7; 1.15: 0.62; 1.16: 0.7 | [8 empty] | 16
| 1 | 4 | 1.03: 0.73057; 1.04: 0.695; 1.09: 0.7; 1.10: 0.8135; 1.11: 0.7; 1.14: 0.7; 1.15: 0.6; 1.16: 0.7 | [8 empty] | 0.7082
| 1 | 5 | 1.03: 0.75107; 1.04: 0.724; 1.09: 0.7; 1.10: 0.8577; 1.11: 0.7; 1.14: 0.7; 1.15: 0.6; 1.16: 0.7 | [8 empty] | 125
| 1 | 6 | 1.03: 0.7568; 1.04: 0.719; 1.09: 0.7; 1.10: 0.8887; 1.11: 0.7; 1.14: 0.7; 1.15: 0.58; 1.16: 0.7
| Mean |  | 1.03: 0.74254; 1.04: 0.70667; 1.09: 0.7; 1.10: 0.83975; 1.11: 0.7; 1.14: 0.7; 1.15: 0.61; 1.16: 0.7
| Median |  | 1.03: 0.7464; 1.04: 0.7055; 1.09: 0.7; 1.10: 0.8325; 1.11: 0.7; 1.14: 0.7; 1.15: 0.605; 1.16: 0.7
| Std Dev. |  | 1.03: 0.014064; 1.04: 0.01401; 1.09: 1.2162e-16; 1.10: 0.029209; 1.11: 1.2162e-16; 1.14: 1.2162e-16; 1.15: 0.023664; 1.16: 1.2162e-16
| Rel.Std.Dev. |  | 1.03: 1.894%; 1.04: 1.982%; 1.09: 1.737e-14%; 1.10: 3.478%; 1.11: 1.737e-14%; 1.14: 1.737e-14%; 1.15: 3.879%; 1.16: 1.737e-14%
| PDM3 |  | 1.03: 4.849%; 1.04: -0.2167%; 1.09: -1.158%; 1.10: 18.58%; 1.11: -1.158%; 1.14: -1.158%; 1.15: -13.87%; 1.16: -1.158%
| Z-Score (Absolute) |  | 1.03: 8.61; 1.04: 1.35; 1.09: 0; 1.10: 28.27; 1.11: 0; 1.14: 0; 1.15: 18.21; 1.16: 0
| Analytical results for Zn in OREAS 284 (Certified Value 77 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 0
|  |  | 0.25g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 1
| 1 | 1 | 71.37 | 80.0244 | 82 | 72.109 | 113 | 81.788 | 72 | 74 | 76.4 | 82 | 85 | 79 | 78 | 81 | 76.6777 | 79 | 80 | 80 | 82.7 | 72.8 | 78 | 74 | 73 | 71.26 | 75.6 | 98.08 | 59 | 74 |  | 1
| 1 | 2 | 71.79 | 78.4089 | 83 | 69.6894 | 109 | 85.2424 | 72 | 73 | 81.4 | 83 | 85 | 79 | 77 | 80 | 79.3938 | 78 | 83 | 78 | 81.6 | 75.5 | 75 | 70 | 78 | 69.54 | 72.8 | 100.56 | 55 | 74 |  | 17
| 1 | 3 | 70.58 | 77.6544 | 82 | 70.744 | 110 | 74.2578 | 72 | 75 | 81 | 83 | 82 | 79 | 78 | 80 | 80.8139 | 79 | 81 | 81 | 80.3 | 71.9 | 76 | 71 | 80 | 71.02 | 73.7 | 95.12 | 56 | 74 |  | 16
| 1 | 4 | 70.84 | 79.239 | 81 | 70.8475 | 109 | 74.0323 | 71 | 73 | 80.6 | 82 | 83 | 78 | 79 | 80 | 77.0951 | 77 | 81 | 80 | 80.2 | 76.1 | 78 | 68 | 77 | 71.83 | 72.9 | 90.8 | 56 | 76 |  | 77.0612
| 1 | 5 | 71.9 | 78.8185 | 81 | 71.5269 | 112 | 75.3526 | 72 | 73 | 81.1 | 83 | 82 | 80 | 79 | 81 | 80.6626 | 77 | 80 | 82 | 79.8 | 74 | 77 | 71 | 81 | 71.78 | 72.4 | 100.16 | 52 | 76 |  | 126
| 1 | 6 | 71.91 | 80.455 | 82 | 70.6537 | 109 | 78.7208 | 72 | 74 | 76.9 | 78 | 79 | 79 | 78 | 83 | 77.2616 | 78 | 81 | 80 | 78.9 | 75.2 | 77 | 68 | 78 | 70.13 | 73 | 96 | 53 | 77
| Mean |  | 71.3983 | 79.1 | 81.8333 | 70.9284 | 110.3333 | 78.2323 | 71.8333 | 73.6667 | 79.5667 | 81.8333 | 82.6667 | 79 | 78.1667 | 80.8333 | 78.6508 | 78 | 81 | 80.1667 | 80.5833 | 74.25 | 76.8333 | 70.3333 | 77.8333 | 70.9267 | 73.4 | 96.7867 | 55.1667 | 75.1667
| Median |  | 71.58 | 79.0288 | 82 | 70.7958 | 109.5 | 77.0367 | 72 | 73.5 | 80.8 | 82.5 | 82.5 | 79 | 78 | 80.5 | 78.3277 | 78 | 81 | 80 | 80.25 | 74.6 | 77 | 70.5 | 78 | 71.14 | 72.95 | 97.04 | 55.5 | 75
| Std Dev. |  | 0.5743 | 1.035 | 0.75277 | 0.82498 | 1.7512 | 4.5553 | 0.40825 | 0.8165 | 2.2792 | 1.9408 | 2.2509 | 0.63246 | 0.75277 | 1.169 | 1.872 | 0.89443 | 1.0954 | 1.3292 | 1.3556 | 1.6477 | 1.169 | 2.2509 | 2.7869 | 0.91886 | 1.1576 | 3.6477 | 2.4833 | 1.3292
| Rel.Std.Dev. |  | 0.8044% | 1.309% | 0.9199% | 1.163% | 1.587% | 5.823% | 0.5683% | 1.108% | 2.864% | 2.372% | 2.723% | 0.8006% | 0.963% | 1.446% | 2.38% | 1.147% | 1.352% | 1.658% | 1.682% | 2.219% | 1.522% | 3.2% | 3.581% | 1.296% | 1.577% | 3.769% | 4.501% | 1.768%
| PDM3 |  | -7.349% | 2.646% | 6.193% | -7.958% | 43.18% | 1.52% | -6.784% | -4.405% | 3.251% | 6.193% | 7.274% | 2.516% | 1.435% | 4.895% | 2.063% | 1.218% | 5.111% | 4.03% | 4.571% | -3.648% | -0.2957% | -8.731% | 1.002% | -7.961% | -4.751% | 25.6% | -28.41% | -2.458%
| Z-Score (Absolute) |  | 1.38 | 0.18 | 0.74 | 1.47 | 6.5 | 0.01 | 1.29 | 0.92 | 0.28 | 0.74 | 0.9 | 0.16 | 0.01 | 0.53 | 0.09 | 0.04 | 0.57 | 0.4 | 0.48 | 0.8 | 0.28 | 1.59 | 0.07 | 1.47 | 0.97 | 3.76 | 4.66 | 0.61
| Analytical results for Zr in OREAS 284 (Certified Value 8.27 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [8 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [8 empty] | 3
| 1 | 1 | 8.152 | 7.2 | 8.7 | 7.5684 | 10 | 8 | 8.5 | 8.4 | 8.2 | 7.3 | 7.6373 | 7.6 | 7.8 | 9.6 | 7.3 | 8.4 | 8.5 | 7.94 | < 0.5 | 9.8 | 8.5 | [8 empty] | 1
| 1 | 2 | 8.0515 | 7.6 | 8.3 | 14.2481 | 10 | 8.1 | 8.5 | 8.3 | 8.1 | 7.6 | 7.5528 | 8.1 | 8.2 | 9.3 | 7.3 | 8.1 | 9.3 | 7.53 | < 0.5 | 9.4 | 8.5 | [8 empty] | 18
| 1 | 3 | 8.2002 | 7.3 | 8.4 | 5.5661 | 10 | 8.3 | 7.8 | 8.1 | 8.1 | 7.3 | 9.1861 | 7.8 | 8 | 9.3 | 7.3 | 8.4 | 9.5 | 7.78 | < 0.5 | 8.8 | 8.7 | [8 empty] | 16
| 1 | 4 | 8.1393 | 7.2 | 8.7 | 5.3611 | 9 | 8.2 | 9.7 | 8.1 | 8.3 | 7.1 | 8.2529 | 7.7 | 8 | 9.5 | 7 | 8.5 | 9.1 | 8.1 | < 0.5 | 9.5 | 8.7 | [8 empty] | 8.274
| 1 | 5 | 8.1781 | 7.3 | 9.4 | 5.1045 | 9 | 8.1 | 8.7 | 8.7 | 8.5 | 7.3 | 9.0215 | 7.8 | 8 | 9.5 | 7.1 | 8.8 | 9.4 | 7.94 | < 0.5 | 8.1 | 8.8 | [8 empty] | 127
| 1 | 6 | 8.3916 | 7.7 | 8.9 | 5.2367 | 9 | 7.3 | 8.2 | 8.4 | 8.6 | 7.3 | 8.2761 | 7.9 | 8 | 9.5 | 7.1 | 8.7 | 9.1 | 8.26 | < 0.5 | 8.6 | 8.8
| Mean |  | 8.1855 | 7.3833 | 8.7333 | 7.1808 | 9.5 | 8 | 8.5667 | 8.3333 | 8.3 | 7.3167 | 8.3211 | 7.8167 | 8 | 9.45 | 7.1833 | 8.4833 | 9.15 | 7.925 |  | 9.0333 | 8.6667
| Median |  | 8.1651 | 7.3 | 8.7 | 5.4636 | 9.5 | 8.1 | 8.5 | 8.35 | 8.25 | 7.3 | 8.2645 | 7.8 | 8 | 9.5 | 7.2 | 8.45 | 9.2 | 7.94 |  | 9.1 | 8.7
| Std Dev. |  | 0.11309 | 0.2137 | 0.39328 | 3.5807 | 0.54772 | 0.35777 | 0.6377 | 0.22509 | 0.20976 | 0.16021 | 0.67872 | 0.17224 | 0.12649 | 0.12247 | 0.13292 | 0.24833 | 0.35637 | 0.25312 |  | 0.64083 | 0.13663
| Rel.Std.Dev. |  | 1.382% | 2.894% | 4.503% | 49.86% | 5.766% | 4.472% | 7.444% | 2.701% | 2.527% | 2.19% | 8.157% | 2.203% | 1.581% | 1.296% | 1.85% | 2.927% | 3.895% | 3.194% |  | 7.094% | 1.576%
| PDM3 |  | -1.07% | -10.76% | 5.552% | -13.21% | 14.82% | -3.311% | 3.537% | 0.7174% | 0.3145% | -11.57% | 0.5696% | -5.527% | -3.311% | 14.21% | -13.18% | 2.53% | 10.59% | -4.218% |  | 9.178% | 4.746%
| Z-Score (Absolute) |  | 0.09 | 1.36 | 0.78 | 1.68 | NA | 0.39 | 0.51 | 0.14 | 0.09 | 1.47 | 0.12 | 0.68 | 0.39 | 1.92 | 1.68 | 0.38 | 1.44 | 0.5 | 12.68 | 1.25 | 0.67
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1020: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1038: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 284 (Indicative Value 0.129 wt.%) | [28 empty] | Indicative
C | Round | Replicate | Lab | BV Geo | [25 empty] | 1
| No | No | 1.12 | Perth | [25 empty] | wt.%
|  |  | IRC | IRC | [25 empty] | 3
3
| 1 | 1 | 0.1 | 0.14 | [25 empty] | 1
| 1 | 2 | 0.12 | 0.14 | [25 empty] | 9
| 1 | 3 | 0.12 | [26 empty] | 16
| 1 | 4 | 0.13 | [26 empty] | 0.12917
| 1 | 5 | 0.13 | [26 empty] | 15
| 1 | 6 | 0.11
| Mean |  | 0.11833 | 0.14
| Median |  | 0.12 | 0.14
| Std Dev. |  | 0.01169 | 0
| Rel.Std.Dev. |  | 9.879% | 0%
| PDM3 |  | -8.387% | 8.387%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 284 (Certified Value 0.257 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.03 | 1.04 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.24916 | 0.25 | 0.28 | 0.26 | 0.27 | 0.28 | 0.13 | 0.24 | 0.25 | 0.26 | 0.24 | 0.22 | 0.27 | 0.26 | 0.27 | 0.27 | 0.27 | 0.22 | 0.26 | 0.253 | 0.24 | 0.26 | 0.29 | 0.24 | 0.244 | 0.14 | 0.25 | 1
| 1 | 2 | 0.25054 | 0.26 | 0.273 | 0.25 | 0.27 | 0.27 | 0.27 | 0.23 | 0.25 | 0.26 | 0.23 | 0.25 | 0.27 | 0.25 | 0.27 | 0.27 | 0.27 | 0.23 | 0.27 | 0.257 | 0.24 | 0.26 | 0.285 | 0.25 | 0.246 | 0.13 | 0.27 | 11
| 1 | 3 | 0.24579 | 0.26 | 0.299 | 0.25 | 0.26 | 0.28 | 0.27 | 0.24 | 0.24 | 0.26 | 0.23 | 0.24 | 0.27 | 0.24867 | 0.27 | 0.26 | 0.27 | 0.22 | 0.26 | 0.253 | 0.25 | 0.26 | 0.285 | 0.25 | 0.254 | 0.14 | 0.28 | 16
| 1 | 4 | 0.25654 | 0.25 | 0.258 | 0.26 | 0.27 | 0.28 | 0.27 | 0.25 | 0.25 | 0.27 | 0.24 | 0.24 | 0.26 | 0.25781 | 0.27 | 0.27 | 0.27 | 0.23 | 0.25 | 0.257 | 0.24 | 0.26 | 0.275 | 0.25 | 0.263 | 0.14 | 0.27 | 0.25716
| 1 | 5 | 0.25194 | 0.25 | 0.264 | 0.26 | 0.26 | 0.28 | 0.26 | 0.25 | 0.25 | 0.26 | 0.25 | 0.24 | 0.26 | 0.24072 | 0.27 | 0.27 | 0.27 | 0.23 | 0.26 | 0.255 | 0.24 | 0.26 | 0.285 | 0.25 | 0.259 | 0.13 | 0.25 | 129
| 1 | 6 | 0.25324 | 0.25 | 0.262 | 0.26 | 0.27 | 0.28 | 0.16 | 0.23 | 0.25 | 0.26 | 0.26 | 0.22 | 0.26 | 0.24099 | 0.27 | 0.27 | 0.27 | 0.23 | 0.26 | 0.255 | 0.24 | 0.25 | 0.28 | 0.26 | 0.265 | 0.13 | 0.24
| Mean |  | 0.2512 | 0.25333 | 0.27267 | 0.25667 | 0.26667 | 0.27833 | 0.22667 | 0.24 | 0.24833 | 0.26167 | 0.24167 | 0.235 | 0.265 | 0.2497 | 0.27 | 0.26833 | 0.27 | 0.22667 | 0.26 | 0.255 | 0.24167 | 0.25833 | 0.28333 | 0.25 | 0.25517 | 0.135 | 0.26
| Median |  | 0.25124 | 0.25 | 0.2685 | 0.26 | 0.27 | 0.28 | 0.265 | 0.24 | 0.25 | 0.26 | 0.24 | 0.24 | 0.265 | 0.24933 | 0.27 | 0.27 | 0.27 | 0.23 | 0.26 | 0.255 | 0.24 | 0.26 | 0.285 | 0.25 | 0.2565 | 0.135 | 0.26
| Std Dev. |  | 0.0036631 | 0.005164 | 0.015175 | 0.005164 | 0.005164 | 0.0040825 | 0.064083 | 0.0089443 | 0.0040825 | 0.0040825 | 0.01169 | 0.012247 | 0.0054772 | 0.0081174 | 0 | 0.0040825 | 0 | 0.005164 | 0.0063246 | 0.0017889 | 0.0040825 | 0.0040825 | 0.005164 | 0.0063246 | 0.0087502 | 0.0054772 | 0.015492
| Rel.Std.Dev. |  | 1.458% | 2.038% | 5.565% | 2.012% | 1.936% | 1.467% | 28.27% | 3.727% | 1.644% | 1.56% | 4.837% | 5.212% | 2.067% | 3.251% | 0% | 1.521% | 0% | 2.278% | 2.433% | 0.7015% | 1.689% | 1.58% | 1.823% | 2.53% | 3.429% | 4.057% | 5.958%
| PDM3 |  | -2.318% | -1.489% | 6.029% | -0.1929% | 3.696% | 8.232% | -11.86% | -6.674% | -3.433% | 1.751% | -6.026% | -8.618% | 3.048% | -2.902% | 4.992% | 4.344% | 4.992% | -11.86% | 1.103% | -0.841% | -6.026% | 0.4552% | 10.18% | -2.785% | -0.7762% | -47.5% | 1.103%
| Z-Score (Absolute) |  | 0.23 | 0.11 | 1.03 | 0.09 | 0.67 | 1.36 | 1.67 | 0.89 | 0.4 | 0.38 | 0.79 | 1.18 | 0.58 | 0.32 | 0.87 | 0.77 | 0.87 | 1.67 | 0.28 | 0.01 | 0.79 | 0.19 | 1.65 | 0.3 | 0 | 7.05 | 0.28
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 284 (Indicative Value 15.17 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 15.16 | 1
| 1 | 2 | 15.17 | 11
| Mean |  | 15.165 | 16
| Median |  | 15.165 | 15.165
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.04663%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 284 (Indicative Value 170 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 170 | 12
| Mean |  | 170 | 16
| Median |  | 170 | 170
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 284 (Indicative Value 1065 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1070 | 1
| 1 | 2 | 1060 | 13
| Mean |  | 1065 | 16
| Median |  | 1065 | 1065
| Std Dev. |  | 7.0711 | 19
| Rel.Std.Dev. |  | 0.664%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 284 (Indicative Value 2.64 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.63 | 1
| 1 | 2 | 2.64 | 14
| Mean |  | 2.635 | 16
| Median |  | 2.635 | 2.635
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.2684%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 284 (Indicative Value 80 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 70 | 15
| Mean |  | 80 | 16
| Median |  | 80 | 80
| Std Dev. |  | 14.1421 | 21
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 284 (Indicative Value 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 16
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 284 (Indicative Value 60 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 70 | 17
| Mean |  | 60 | 16
| Median |  | 60 | 60
| Std Dev. |  | 14.1421 | 23
| Rel.Std.Dev. |  | 23.57%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 284 (Indicative Value 35 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 30 | 18
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 7.0711 | 24
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 284 (Indicative Value 3.24 wt.%) |  |  | Indicative
Fe | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.25 | 1
| 1 | 2 | 3.22 | 11
| Mean |  | 3.235 | 16
| Median |  | 3.235 | 3.235
| Std Dev. |  | 0.021213 | 17
| Rel.Std.Dev. |  | 0.6557%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 284 (Indicative Value 3.55 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.55 | 1
| 1 | 2 | 3.55 | 12
| Mean |  | 3.55 | 16
| Median |  | 3.55 | 3.55
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 284 (Indicative Value 1.49 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.49 | 1
| 1 | 2 | 1.48 | 13
| Mean |  | 1.485 | 16
| Median |  | 1.485 | 1.485
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.4762%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 284 (Indicative Value 0.051 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.051 | 1
| 1 | 2 | 0.05 | 14
| Mean |  | 0.0505 | 16
| Median |  | 0.0505 | 0.0505
| Std Dev. |  | 0.00070711 | 20
| Rel.Std.Dev. |  | 1.4%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 284 (Indicative Value 2.76 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.76 | 1
| 1 | 2 | 2.76 | 15
| Mean |  | 2.76 | 16
| Median |  | 2.76 | 2.76
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 284 (Indicative Value 25 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 20 | 16
| Mean |  | 25 | 16
| Median |  | 25 | 25
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 284 (Indicative Value 0.086 wt.%) |  |  | Indicative
P | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.086 | 1
| 1 | 2 | 0.086 | 17
| Mean |  | 0.086 | 16
| Median |  | 0.086 | 0.086
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 284 (Indicative Value 20 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 18
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 284 (Indicative Value 0.262 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.262 | 1
| 1 | 2 | 0.262 | 11
| Mean |  | 0.262 | 16
| Median |  | 0.262 | 0.262
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 284 (Indicative Value 67.52 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.57 | 1
| 1 | 2 | 67.46 | 12
| Mean |  | 67.515 | 16
| Median |  | 67.515 | 67.515
| Std Dev. |  | 0.077782 | 18
| Rel.Std.Dev. |  | 0.1152%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 284 (Indicative Value < 10 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 13
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 19
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 284 (Indicative Value 155 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 140 | 14
| Mean |  | 155 | 16
| Median |  | 155 | 155
| Std Dev. |  | 21.2132 | 20
| Rel.Std.Dev. |  | 13.69%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 284 (Indicative Value 0.633 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.633 | 1
| 1 | 2 | 0.633 | 15
| Mean |  | 0.633 | 16
| Median |  | 0.633 | 0.633
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 284 (Indicative Value 60 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 60 | 16
| Mean |  | 60 | 16
| Median |  | 60 | 60
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 284 (Indicative Value 80 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 80 | 17
| Mean |  | 80 | 16
| Median |  | 80 | 80
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 284 (Indicative Value 240 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 240 | 1
| 1 | 2 | 240 | 18
| Mean |  | 240 | 16
| Median |  | 240 | 240
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 284 (Indicative Value 1.17 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.18 | 1
| 1 | 2 | 1.16 | 20
| Mean |  | 1.17 | 16
| Median |  | 1.17 | 1.17
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 1.209%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 284 (Indicative Value 0.1 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.1 | 22
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.1
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 284 (Indicative Value 165 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 165 | 1
| 1 | 2 | 165 | 12
| Mean |  | 165 | 16
| Median |  | 165 | 165
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 284 (Indicative Value 999 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1000 | 1
| 1 | 2 | 997 | 13
| Mean |  | 998.5 | 16
| Median |  | 998.5 | 998.5
| Std Dev. |  | 2.1213 | 30
| Rel.Std.Dev. |  | 0.2125%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 284 (Indicative Value 2.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.6 | 25
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 284 (Indicative Value 0.6 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.62 | 1
| 1 | 2 | 0.58 | 26
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0.028284 | 32
| Rel.Std.Dev. |  | 4.714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 284 (Indicative Value 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.1 | 27
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.1
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 284 (Indicative Value 79 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 78.3 | 28
| Mean |  | 79.15 | 16
| Median |  | 79.15 | 79.15
| Std Dev. |  | 1.2021 | 34
| Rel.Std.Dev. |  | 1.519%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 284 (Indicative Value 12.3 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12 | 1
| 1 | 2 | 12.5 | 16
| Mean |  | 12.25 | 16
| Median |  | 12.25 | 12.25
| Std Dev. |  | 0.35355 | 35
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 284 (Indicative Value 51 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 52 | 1
| 1 | 2 | 50 | 17
| Mean |  | 51 | 16
| Median |  | 51 | 51
| Std Dev. |  | 1.4142 | 36
| Rel.Std.Dev. |  | 2.773%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 284 (Indicative Value 11.3 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.3 | 1
| 1 | 2 | 11.2 | 31
| Mean |  | 11.25 | 16
| Median |  | 11.25 | 11.25
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.6285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 284 (Indicative Value 28 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28 | 1
| 1 | 2 | 28 | 18
| Mean |  | 28 | 16
| Median |  | 28 | 28
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 284 (Indicative Value 6.07 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6 | 1
| 1 | 2 | 6.13 | 33
| Mean |  | 6.065 | 16
| Median |  | 6.065 | 6.065
| Std Dev. |  | 0.091924 | 39
| Rel.Std.Dev. |  | 1.516%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 284 (Indicative Value 3.12 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.12 | 1
| 1 | 2 | 3.12 | 34
| Mean |  | 3.12 | 16
| Median |  | 3.12 | 3.12
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 284 (Indicative Value 1.46 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.48 | 1
| 1 | 2 | 1.43 | 35
| Mean |  | 1.455 | 16
| Median |  | 1.455 | 1.455
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 2.43%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 284 (Indicative Value 20 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.9 | 1
| 1 | 2 | 20.1 | 36
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0.14142 | 42
| Rel.Std.Dev. |  | 0.7071%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 284 (Indicative Value 7.07 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.08 | 1
| 1 | 2 | 7.05 | 37
| Mean |  | 7.065 | 16
| Median |  | 7.065 | 7.065
| Std Dev. |  | 0.021213 | 43
| Rel.Std.Dev. |  | 0.3003%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 284 (Indicative Value 1.38 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.4 | 2
| Mean |  | 1.375 | 16
| Median |  | 1.375 | 1.375
| Std Dev. |  | 0.035355 | 44
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 284 (Indicative Value 7.06 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.14 | 1
| 1 | 2 | 6.97 | 22
| Mean |  | 7.055 | 16
| Median |  | 7.055 | 7.055
| Std Dev. |  | 0.12021 | 28
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 284 (Indicative Value 1.15 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.15 | 1
| 1 | 2 | 1.14 | 23
| Mean |  | 1.145 | 16
| Median |  | 1.145 | 1.145
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.6176%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 284 (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 24
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 284 (Indicative Value 39.7 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39.9 | 1
| 1 | 2 | 39.4 | 25
| Mean |  | 39.65 | 16
| Median |  | 39.65 | 39.65
| Std Dev. |  | 0.35355 | 31
| Rel.Std.Dev. |  | 0.8917%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 284 (Indicative Value 0.39 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.4 | 26
| Mean |  | 0.39 | 16
| Median |  | 0.39 | 0.39
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 284 (Indicative Value 0.038 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0373 | 1
| 1 | 2 | 0.038 | 27
| Mean |  | 0.03765 | 16
| Median |  | 0.03765 | 0.03765
| Std Dev. |  | 0.00049497 | 33
| Rel.Std.Dev. |  | 1.315%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 284 (Indicative Value 2.4 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.4 | 28
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 284 (Indicative Value 13 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.9 | 1
| 1 | 2 | 13.1 | 29
| Mean |  | 13 | 16
| Median |  | 13 | 13
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 1.088%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 284 (Indicative Value 37.5 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 37.4 | 1
| 1 | 2 | 37.6 | 6
| Mean |  | 37.5 | 16
| Median |  | 37.5 | 37.5
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 0.3771%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 284 (Indicative Value 19 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 20 | 16
| Mean |  | 19 | 16
| Median |  | 19 | 19
| Std Dev. |  | 1.4142 | 37
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 284 (Indicative Value 23 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23 | 1
| 1 | 2 | 23 | 18
| Mean |  | 23 | 16
| Median |  | 23 | 23
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 284 (Indicative Value 9.76 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.74 | 1
| 1 | 2 | 9.77 | 4
| Mean |  | 9.755 | 16
| Median |  | 9.755 | 9.755
| Std Dev. |  | 0.021213 | 39
| Rel.Std.Dev. |  | 0.2175%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 284 (Indicative Value 165 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 165 | 34
| Mean |  | 164.5 | 16
| Median |  | 164.5 | 164.5
| Std Dev. |  | 0.70711 | 40
| Rel.Std.Dev. |  | 0.4299%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 284 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 2
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 284 (Indicative Value 0.4 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 36
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 284 (Indicative Value 8.3 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.2 | 1
| 1 | 2 | 8.4 | 37
| Mean |  | 8.3 | 16
| Median |  | 8.3 | 8.3
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 284 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 38
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 284 (Indicative Value 7.88 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.9 | 1
| 1 | 2 | 7.86 | 4
| Mean |  | 7.88 | 16
| Median |  | 7.88 | 7.88
| Std Dev. |  | 0.028284 | 28
| Rel.Std.Dev. |  | 0.3589%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 284 (Indicative Value 4.5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.6 | 13
| Mean |  | 4.5 | 16
| Median |  | 4.5 | 4.5
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 284 (Indicative Value 157 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 157 | 1
| 1 | 2 | 157 | 14
| Mean |  | 157 | 16
| Median |  | 157 | 157
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 284 (Indicative Value 1.11 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.11 | 1
| 1 | 2 | 1.1 | 25
| Mean |  | 1.105 | 16
| Median |  | 1.105 | 1.105
| Std Dev. |  | 0.0070711 | 31
| Rel.Std.Dev. |  | 0.6399%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 284 (Indicative Value 1.11 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.12 | 1
| 1 | 2 | 1.09 | 26
| Mean |  | 1.105 | 16
| Median |  | 1.105 | 1.105
| Std Dev. |  | 0.021213 | 32
| Rel.Std.Dev. |  | 1.92%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 284 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 284 (Indicative Value 15.3 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.4 | 1
| 1 | 2 | 15.2 | 28
| Mean |  | 15.3 | 16
| Median |  | 15.3 | 15.3
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 0.9243%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 284 (Indicative Value 0.374 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.374 | 1
| 1 | 2 | 0.374 | 29
| Mean |  | 0.374 | 16
| Median |  | 0.374 | 0.374
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 284 (Indicative Value 0.8 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.8 | 30
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 284 (Indicative Value 0.44 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.43 | 5
| Mean |  | 0.44 | 16
| Median |  | 0.44 | 0.44
| Std Dev. |  | 0.014142 | 37
| Rel.Std.Dev. |  | 3.214%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 284 (Indicative Value 4.7 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.73 | 1
| 1 | 2 | 4.66 | 32
| Mean |  | 4.695 | 16
| Median |  | 4.695 | 4.695
| Std Dev. |  | 0.049497 | 38
| Rel.Std.Dev. |  | 1.054%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 284 (Indicative Value 64 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 64 | 1
| 1 | 2 | 64.6 | 16
| Mean |  | 64.3 | 16
| Median |  | 64.3 | 64.3
| Std Dev. |  | 0.42426 | 39
| Rel.Std.Dev. |  | 0.6598%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 284 (Indicative Value 17.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.5 | 1
| 1 | 2 | 17.5 | 34
| Mean |  | 17.5 | 16
| Median |  | 17.5 | 17.5
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 284 (Indicative Value 30.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.8 | 1
| 1 | 2 | 30.4 | 35
| Mean |  | 30.6 | 16
| Median |  | 30.6 | 30.6
| Std Dev. |  | 0.28284 | 41
| Rel.Std.Dev. |  | 0.9243%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 284 (Indicative Value 2.86 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.82 | 1
| 1 | 2 | 2.9 | 36
| Mean |  | 2.86 | 16
| Median |  | 2.86 | 2.86
| Std Dev. |  | 0.056569 | 42
| Rel.Std.Dev. |  | 1.978%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 284 (Indicative Value 75 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75 | 1
| 1 | 2 | 75 | 17
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 284 (Indicative Value 248 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 250 | 1
| 1 | 2 | 245 | 18
| Mean |  | 247.5 | 16
| Median |  | 247.5 | 247.5
| Std Dev. |  | 3.5355 | 44
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 284 (Indicative Value 0.223 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.23 | 1
| 1 | 2 | 0.23 | 22
| 1 | 3 | 0.22 | 16
| 1 | 4 | 0.21 | 0.22333
| 1 | 5 | 0.23 | 46
| 1 | 6 | 0.22
| Mean |  | 0.22333
| Median |  | 0.225
| Std Dev. |  | 0.008165
| Rel.Std.Dev. |  | 3.656%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 284 (Indicative Value 15.85 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 15.7 | 1
| 1 | 2 | 15.9 | 11
| 1 | 3 | 15.5 | 16
| 1 | 4 | 15.6 | 15.85
| 1 | 5 | 16 | 47
| 1 | 6 | 16.4
| Mean |  | 15.85
| Median |  | 15.8
| Std Dev. |  | 0.32711
| Rel.Std.Dev. |  | 2.064%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 284 (Indicative Value 1067 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 1090 | 1
| 1 | 2 | 1080 | 13
| 1 | 3 | 1040 | 16
| 1 | 4 | 1050 | 1066.6667
| 1 | 5 | 1080 | 48
| 1 | 6 | 1060
| Mean |  | 1066.6667
| Median |  | 1070
| Std Dev. |  | 19.6638
| Rel.Std.Dev. |  | 1.843%
| PDM3 |  | -2.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 284 (Indicative Value 3 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.9 | 1
| 1 | 2 | 3.1 | 25
| 1 | 3 | 3 | 16
| 1 | 4 | 3 | 3
| 1 | 5 | 3 | 49
| 1 | 6 | 3
| Mean |  | 3
| Median |  | 3
| Std Dev. |  | 0.063246
| Rel.Std.Dev. |  | 2.108%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 284 (Indicative Value 0.6 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.5 | 26
| 1 | 3 | 0.6 | 16
| 1 | 4 | 0.7 | 0.6
| 1 | 5 | 0.6 | 50
| 1 | 6 | 0.5
| Mean |  | 0.6
| Median |  | 0.6
| Std Dev. |  | 0.089443
| Rel.Std.Dev. |  | 14.91%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 284 (Indicative Value 2.64 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 2.62 | 1
| 1 | 2 | 2.67 | 14
| 1 | 3 | 2.58 | 16
| 1 | 4 | 2.59 | 2.6417
| 1 | 5 | 2.66 | 51
| 1 | 6 | 2.73
| Mean |  | 2.6417
| Median |  | 2.64
| Std Dev. |  | 0.056362
| Rel.Std.Dev. |  | 2.134%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 284 (Indicative Value 0.23 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.3 | 27
| 1 | 3 | 0.2 | 16
| 1 | 4 | 0.2 | 0.23333
| 1 | 5 | 0.3 | 52
| 1 | 6 | 0.2
| Mean |  | 0.23333
| Median |  | 0.2
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 22.13%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 284 (Indicative Value 78 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 78 | 1
| 1 | 2 | 79 | 28
| 1 | 3 | 77 | 16
| 1 | 4 | 76 | 77.6667
| 1 | 5 | 78 | 53
| 1 | 6 | 78
| Mean |  | 77.6667
| Median |  | 78
| Std Dev. |  | 1.0328
| Rel.Std.Dev. |  | 1.33%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 284 (Indicative Value 12.7 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 12.5 | 1
| 1 | 2 | 13 | 16
| 1 | 3 | 12.4 | 16
| 1 | 4 | 12.5 | 12.65
| 1 | 5 | 12.8 | 54
| 1 | 6 | 12.7
| Mean |  | 12.65
| Median |  | 12.6
| Std Dev. |  | 0.22583
| Rel.Std.Dev. |  | 1.785%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 284 (Indicative Value 44.3 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 45 | 17
| 1 | 3 | 46 | 16
| 1 | 4 | 42 | 44.3333
| 1 | 5 | 40 | 55
| 1 | 6 | 43
| Mean |  | 44.3333
| Median |  | 44
| Std Dev. |  | 3.5024
| Rel.Std.Dev. |  | 7.9%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 284 (Indicative Value 9.33 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 9.4 | 1
| 1 | 2 | 9.5 | 31
| 1 | 3 | 9.3 | 16
| 1 | 4 | 9.2 | 9.3333
| 1 | 5 | 9.3 | 56
| 1 | 6 | 9.3
| Mean |  | 9.3333
| Median |  | 9.3
| Std Dev. |  | 0.10328
| Rel.Std.Dev. |  | 1.107%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 284 (Indicative Value 27.5 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 31.2 | 1
| 1 | 2 | 29.5 | 18
| 1 | 3 | 25.4 | 16
| 1 | 4 | 26.8 | 27.4833
| 1 | 5 | 26.2 | 57
| 1 | 6 | 25.8
| Mean |  | 27.4833
| Median |  | 26.5
| Std Dev. |  | 2.331
| Rel.Std.Dev. |  | 8.482%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 284 (Indicative Value 3.61 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.58 | 1
| 1 | 2 | 3.73 | 33
| 1 | 3 | 3.59 | 16
| 1 | 4 | 3.55 | 3.6067
| 1 | 5 | 3.63 | 58
| 1 | 6 | 3.56
| Mean |  | 3.6067
| Median |  | 3.585
| Std Dev. |  | 0.066533
| Rel.Std.Dev. |  | 1.845%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 284 (Indicative Value 1.4 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.42 | 1
| 1 | 2 | 1.48 | 34
| 1 | 3 | 1.38 | 16
| 1 | 4 | 1.35 | 1.4
| 1 | 5 | 1.4 | 59
| 1 | 6 | 1.37
| Mean |  | 1.4
| Median |  | 1.39
| Std Dev. |  | 0.046043
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 284 (Indicative Value 1.6 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.62 | 1
| 1 | 2 | 1.66 | 35
| 1 | 3 | 1.57 | 16
| 1 | 4 | 1.56 | 1.6033
| 1 | 5 | 1.62 | 60
| 1 | 6 | 1.59
| Mean |  | 1.6033
| Median |  | 1.605
| Std Dev. |  | 0.037238
| Rel.Std.Dev. |  | 2.323%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 284 (Indicative Value 4.68 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 4.66 | 1
| 1 | 2 | 4.72 | 55
| 1 | 3 | 4.59 | 16
| 1 | 4 | 4.58 | 4.6783
| 1 | 5 | 4.73 | 61
| 1 | 6 | 4.79
| Mean |  | 4.6783
| Median |  | 4.69
| Std Dev. |  | 0.083287
| Rel.Std.Dev. |  | 1.78%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 284 (Indicative Value 18.6 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 18.3 | 1
| 1 | 2 | 18.9 | 36
| 1 | 3 | 18.3 | 16
| 1 | 4 | 18.2 | 18.5833
| 1 | 5 | 18.9 | 62
| 1 | 6 | 18.9
| Mean |  | 18.5833
| Median |  | 18.6
| Std Dev. |  | 0.34881
| Rel.Std.Dev. |  | 1.877%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 284 (Indicative Value 4.82 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 5 | 37
| 1 | 3 | 4.8 | 16
| 1 | 4 | 4.7 | 4.8167
| 1 | 5 | 4.8 | 46
| 1 | 6 | 4.8
| Mean |  | 4.8167
| Median |  | 4.8
| Std Dev. |  | 0.098319
| Rel.Std.Dev. |  | 2.041%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 284 (Indicative Value 2.43 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.5 | 22
| 1 | 3 | 2.4 | 16
| 1 | 4 | 2.4 | 2.4333
| 1 | 5 | 2.4 | 47
| 1 | 6 | 2.3
| Mean |  | 2.4333
| Median |  | 2.4
| Std Dev. |  | 0.10328
| Rel.Std.Dev. |  | 4.244%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 284 (Indicative Value 0.51 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.52 | 1
| 1 | 2 | 0.53 | 23
| 1 | 3 | 0.5 | 16
| 1 | 4 | 0.5 | 0.51
| 1 | 5 | 0.51 | 48
| 1 | 6 | 0.5
| Mean |  | 0.51
| Median |  | 0.505
| Std Dev. |  | 0.012649
| Rel.Std.Dev. |  | 2.48%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 284 (Indicative Value 3.69 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.67 | 1
| 1 | 2 | 3.74 | 12
| 1 | 3 | 3.6 | 16
| 1 | 4 | 3.59 | 3.6867
| 1 | 5 | 3.7 | 49
| 1 | 6 | 3.82
| Mean |  | 3.6867
| Median |  | 3.685
| Std Dev. |  | 0.087101
| Rel.Std.Dev. |  | 2.363%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 284 (Indicative Value 38.8 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 38 | 1
| 1 | 2 | 39 | 25
| 1 | 3 | 38 | 16
| 1 | 4 | 38 | 38.8333
| 1 | 5 | 40 | 50
| 1 | 6 | 40
| Mean |  | 38.8333
| Median |  | 38.5
| Std Dev. |  | 0.98319
| Rel.Std.Dev. |  | 2.532%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 284 (Indicative Value 53 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 54 | 1
| 1 | 2 | 54 | 45
| 1 | 3 | 51 | 16
| 1 | 4 | 52 | 52.6667
| 1 | 5 | 53 | 51
| 1 | 6 | 52
| Mean |  | 52.6667
| Median |  | 52.5
| Std Dev. |  | 1.2111
| Rel.Std.Dev. |  | 2.299%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 284 (Indicative Value 1.53 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.51 | 1
| 1 | 2 | 1.55 | 13
| 1 | 3 | 1.5 | 16
| 1 | 4 | 1.51 | 1.5317
| 1 | 5 | 1.55 | 52
| 1 | 6 | 1.57
| Mean |  | 1.5317
| Median |  | 1.53
| Std Dev. |  | 0.028577
| Rel.Std.Dev. |  | 1.866%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 284 (Indicative Value 0.049 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.048 | 1
| 1 | 2 | 0.049 | 14
| 1 | 3 | 0.048 | 16
| 1 | 4 | 0.048 | 0.048833
| 1 | 5 | 0.05 | 53
| 1 | 6 | 0.05
| Mean |  | 0.048833
| Median |  | 0.0485
| Std Dev. |  | 0.00098319
| Rel.Std.Dev. |  | 2.013%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 284 (Indicative Value 2.91 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.98 | 1
| 1 | 2 | 2.57 | 28
| 1 | 3 | 2.96 | 16
| 1 | 4 | 2.57 | 2.9117
| 1 | 5 | 2.57 | 54
| 1 | 6 | 3.82
| Mean |  | 2.9117
| Median |  | 2.765
| Std Dev. |  | 0.48627
| Rel.Std.Dev. |  | 16.7%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 284 (Indicative Value 2.82 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.85 | 15
| 1 | 3 | 2.75 | 16
| 1 | 4 | 2.73 | 2.815
| 1 | 5 | 2.83 | 55
| 1 | 6 | 2.93
| Mean |  | 2.815
| Median |  | 2.815
| Std Dev. |  | 0.072595
| Rel.Std.Dev. |  | 2.579%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 284 (Indicative Value 13.3 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13.3 | 1
| 1 | 2 | 13.6 | 29
| 1 | 3 | 13.1 | 16
| 1 | 4 | 13.1 | 13.2667
| 1 | 5 | 13.3 | 56
| 1 | 6 | 13.2
| Mean |  | 13.2667
| Median |  | 13.25
| Std Dev. |  | 0.18619
| Rel.Std.Dev. |  | 1.403%
| PDM3 |  | -2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 284 (Indicative Value 34 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 33.6 | 1
| 1 | 2 | 35 | 6
| 1 | 3 | 33.7 | 16
| 1 | 4 | 33.9 | 33.9833
| 1 | 5 | 33.7 | 57
| 1 | 6 | 34
| Mean |  | 33.9833
| Median |  | 33.8
| Std Dev. |  | 0.51929
| Rel.Std.Dev. |  | 1.528%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 284 (Indicative Value 22.8 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 23.5 | 1
| 1 | 2 | 24.1 | 16
| 1 | 3 | 21.8 | 16
| 1 | 4 | 22.8 | 22.7833
| 1 | 5 | 22.6 | 58
| 1 | 6 | 21.9
| Mean |  | 22.7833
| Median |  | 22.7
| Std Dev. |  | 0.89759
| Rel.Std.Dev. |  | 3.94%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 284 (Indicative Value 0.201 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.204 | 1
| 1 | 2 | 0.2 | 53
| 1 | 3 | 0.197 | 16
| 1 | 4 | 0.2 | 0.20067
| 1 | 5 | 0.201 | 59
| 1 | 6 | 0.202
| Mean |  | 0.20067
| Median |  | 0.2005
| Std Dev. |  | 0.0023381
| Rel.Std.Dev. |  | 1.165%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 284 (Indicative Value 23.9 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 24.4 | 1
| 1 | 2 | 24.3 | 18
| 1 | 3 | 23.6 | 16
| 1 | 4 | 23.3 | 23.85
| 1 | 5 | 23.7 | 60
| 1 | 6 | 23.8
| Mean |  | 23.85
| Median |  | 23.75
| Std Dev. |  | 0.42308
| Rel.Std.Dev. |  | 1.774%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 284 (Indicative Value 8.57 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 8.5 | 1
| 1 | 2 | 8.8 | 4
| 1 | 3 | 8.6 | 16
| 1 | 4 | 8.4 | 8.5667
| 1 | 5 | 8.6 | 61
| 1 | 6 | 8.5
| Mean |  | 8.5667
| Median |  | 8.55
| Std Dev. |  | 0.13663
| Rel.Std.Dev. |  | 1.595%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 284 (Indicative Value 194 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 191 | 1
| 1 | 2 | 196 | 34
| 1 | 3 | 192 | 16
| 1 | 4 | 191 | 194
| 1 | 5 | 196 | 62
| 1 | 6 | 198
| Mean |  | 194
| Median |  | 194
| Std Dev. |  | 3.0332
| Rel.Std.Dev. |  | 1.563%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 284 (Indicative Value 0.281 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.291 | 1
| 1 | 2 | 0.292 | 11
| 1 | 3 | 0.263 | 16
| 1 | 4 | 0.268 | 0.2805
| 1 | 5 | 0.283 | 46
| 1 | 6 | 0.286
| Mean |  | 0.2805
| Median |  | 0.2845
| Std Dev. |  | 0.012178
| Rel.Std.Dev. |  | 4.341%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 284 (Indicative Value 8.4 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 8.3 | 1
| 1 | 2 | 8.7 | 37
| 1 | 3 | 8.2 | 16
| 1 | 4 | 8.3 | 8.4
| 1 | 5 | 8.5 | 47
| 1 | 6 | 8.4
| Mean |  | 8.4
| Median |  | 8.35
| Std Dev. |  | 0.17889
| Rel.Std.Dev. |  | 2.13%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 284 (Indicative Value 6.67 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.7 | 1
| 1 | 2 | 6.9 | 4
| 1 | 3 | 6.6 | 16
| 1 | 4 | 6.6 | 6.6667
| 1 | 5 | 6.6 | 48
| 1 | 6 | 6.6
| Mean |  | 6.6667
| Median |  | 6.6
| Std Dev. |  | 0.12111
| Rel.Std.Dev. |  | 1.817%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 284 (Indicative Value 4.21 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.11 | 1
| 1 | 2 | 4.36 | 13
| 1 | 3 | 4.24 | 16
| 1 | 4 | 4.12 | 4.2133
| 1 | 5 | 4.23 | 49
| 1 | 6 | 4.22
| Mean |  | 4.2133
| Median |  | 4.225
| Std Dev. |  | 0.091579
| Rel.Std.Dev. |  | 2.174%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 284 (Indicative Value 165 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 165 | 1
| 1 | 2 | 167 | 14
| 1 | 3 | 160 | 16
| 1 | 4 | 160 | 165
| 1 | 5 | 166 | 50
| 1 | 6 | 172
| Mean |  | 165
| Median |  | 165.5
| Std Dev. |  | 4.5607
| Rel.Std.Dev. |  | 2.764%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 284 (Indicative Value 0.93 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.97 | 25
| 1 | 3 | 0.92 | 16
| 1 | 4 | 0.92 | 0.93167
| 1 | 5 | 0.92 | 51
| 1 | 6 | 0.91
| Mean |  | 0.93167
| Median |  | 0.92
| Std Dev. |  | 0.023166
| Rel.Std.Dev. |  | 2.487%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 284 (Indicative Value 0.65 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.66 | 1
| 1 | 2 | 0.68 | 26
| 1 | 3 | 0.64 | 16
| 1 | 4 | 0.64 | 0.65333
| 1 | 5 | 0.65 | 52
| 1 | 6 | 0.65
| Mean |  | 0.65333
| Median |  | 0.65
| Std Dev. |  | 0.015055
| Rel.Std.Dev. |  | 2.304%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 284 (Indicative Value 15.5 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 15.9 | 1
| 1 | 2 | 15.8 | 28
| 1 | 3 | 15.2 | 16
| 1 | 4 | 15.2 | 15.45
| 1 | 5 | 15.3 | 53
| 1 | 6 | 15.3
| Mean |  | 15.45
| Median |  | 15.3
| Std Dev. |  | 0.31464
| Rel.Std.Dev. |  | 2.037%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 284 (Indicative Value 0.66 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.659 | 1
| 1 | 2 | 0.667 | 15
| 1 | 3 | 0.648 | 16
| 1 | 4 | 0.642 | 0.65967
| 1 | 5 | 0.664 | 54
| 1 | 6 | 0.678
| Mean |  | 0.65967
| Median |  | 0.6615
| Std Dev. |  | 0.013095
| Rel.Std.Dev. |  | 1.985%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 284 (Indicative Value 3.81 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.38 | 1
| 1 | 2 | 3.78 | 32
| 1 | 3 | 3.85 | 16
| 1 | 4 | 3.42 | 3.8117
| 1 | 5 | 3.74 | 55
| 1 | 6 | 3.7
| Mean |  | 3.8117
| Median |  | 3.76
| Std Dev. |  | 0.31511
| Rel.Std.Dev. |  | 8.267%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 284 (Indicative Value 66 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 64.9 | 1
| 1 | 2 | 67.7 | 16
| 1 | 3 | 63.4 | 16
| 1 | 4 | 64.6 | 66
| 1 | 5 | 66.8 | 56
| 1 | 6 | 68.6
| Mean |  | 66
| Median |  | 65.85
| Std Dev. |  | 2.011
| Rel.Std.Dev. |  | 3.047%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 284 (Indicative Value 18.2 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 19.3 | 1
| 1 | 2 | 20 | 34
| 1 | 3 | 17.4 | 16
| 1 | 4 | 17.4 | 18.2167
| 1 | 5 | 17.4 | 57
| 1 | 6 | 17.8
| Mean |  | 18.2167
| Median |  | 17.6
| Std Dev. |  | 1.1427
| Rel.Std.Dev. |  | 6.273%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 284 (Indicative Value 13.7 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13.5 | 1
| 1 | 2 | 14.3 | 35
| 1 | 3 | 13.5 | 16
| 1 | 4 | 13.6 | 13.7333
| 1 | 5 | 13.9 | 58
| 1 | 6 | 13.6
| Mean |  | 13.7333
| Median |  | 13.6
| Std Dev. |  | 0.31411
| Rel.Std.Dev. |  | 2.287%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 284 (Indicative Value 1.01 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.07 | 1
| 1 | 2 | 1.06 | 36
| 1 | 3 | 0.98 | 16
| 1 | 4 | 0.99 | 1.0133
| 1 | 5 | 1.01 | 59
| 1 | 6 | 0.97
| Mean |  | 1.0133
| Median |  | 1
| Std Dev. |  | 0.042269
| Rel.Std.Dev. |  | 4.171%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 284 (Indicative Value 82 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 81 | 1
| 1 | 2 | 83 | 17
| 1 | 3 | 81 | 16
| 1 | 4 | 82 | 81.8333
| 1 | 5 | 83 | 60
| 1 | 6 | 81
| Mean |  | 81.8333
| Median |  | 81.5
| Std Dev. |  | 0.98319
| Rel.Std.Dev. |  | 1.201%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 284 (Indicative Value 55 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 56 | 1
| 1 | 2 | 56 | 18
| 1 | 3 | 54 | 16
| 1 | 4 | 54 | 55.3333
| 1 | 5 | 56 | 61
| 1 | 6 | 56
| Mean |  | 55.3333
| Median |  | 56
| Std Dev. |  | 1.0328
| Rel.Std.Dev. |  | 1.866%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.