## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 233b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 1.075 | 1.0601 | 1.0899 | 1.0675 | 1.0825
PhotonAssay (recommended gross mass* 372±15 g)
Au, Gold (ppm) | 1.0698 | 1.0578 | 1.0819 | 1.0678 | 1.0719
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.0124 | 0.99784 | 1.0269 | 1.0046 | 1.0201
Cyanide Leach
Au, Gold (ppm) | 1.0053 | 0.98259 | 1.0281 | 1.0026 | 1.0081
4-Acid Digestion
Ag, Silver (ppm) | 0.31065 | 0.2901 | 0.33119 | 0.29023 | 0.33106
Al, Aluminium (wt.%) | 6.7985 | 6.597 | 7.0001 | 6.6968 | 6.9003
As, Arsenic (ppm) | 59.856 | 57.2995 | 62.4124 | 57.7614 | 61.9505
Ba, Barium (ppm) | 270.0018 | 260.9646 | 279.0389 | 261.5372 | 278.4663
Be, Beryllium (ppm) | 0.39138 | 0.36098 | 0.42178 | 0.36424 | 0.41852
Bi, Bismuth (ppm) | 0.045568 | 0.034309 | 0.056828 | IND | IND
Ca, Calcium (wt.%) | 6.3251 | 6.1641 | 6.4861 | 6.2094 | 6.4408
Cd, Cadmium (ppm) | 0.58174 | 0.53709 | 0.6264 | 0.55251 | 0.61098
Ce, Cerium (ppm) | 12.5654 | 11.9672 | 13.1637 | 12.2183 | 12.9126
Co, Cobalt (ppm) | 43.5587 | 41.8729 | 45.2445 | 42.3326 | 44.7848
Cr, Chromium (ppm) | 146.0186 | 138.6199 | 153.4173 | 141.8018 | 150.2354
Cs, Caesium (ppm) | 0.89133 | 0.83383 | 0.94883 | 0.85823 | 0.92443
Cu, Copper (ppm) | 160.8734 | 155.4953 | 166.2514 | 158.2 | 163.5468
Dy, Dysprosium (ppm) | 3.6044 | 3.3708 | 3.8379 | 3.48 | 3.7288
Er, Erbium (ppm) | 2.1847 | 2.0488 | 2.3206 | 2.1099 | 2.2595
Eu, Europium (ppm) | 0.89567 | 0.82949 | 0.96186 | 0.86125 | 0.93009
Fe, Iron (wt.%) | 7.8431 | 7.61 | 8.0761 | 7.7241 | 7.9621
Ga, Gallium (ppm) | 15.0173 | 14.227 | 15.8075 | 14.5182 | 15.5163
Gd, Gadolinium (ppm) | 2.9573 | 2.6849 | 3.2296 | 2.8395 | 3.075
Ge, Germanium (ppm) | 0.08275 | 0.049856 | 0.11564 | IND | IND
Hf, Hafnium (ppm) | 1.5976 | 1.4845 | 1.7107 | 1.5251 | 1.6701
Ho, Holmium (ppm) | 0.74299 | 0.68693 | 0.79906 | 0.71014 | 0.77585
In, Indium (ppm) | 0.073004 | 0.064529 | 0.08148 | 0.063683 | 0.082326
K, Potassium (wt.%) | 0.54135 | 0.52259 | 0.56011 | 0.52783 | 0.55488
La, Lanthanum (ppm) | 5.301 | 5.0642 | 5.5379 | 5.165 | 5.437
Li, Lithium (ppm) | 12.192 | 11.4794 | 12.9045 | 11.7207 | 12.6632
Lu, Lutetium (ppm) | 0.32391 | 0.29016 | 0.35767 | 0.31171 | 0.33612
Mg, Magnesium (wt.%) | 3.8702 | 3.7579 | 3.9825 | 3.8008 | 3.9397
Mn, Manganese (wt.%) | 0.14715 | 0.14359 | 0.15071 | 0.14435 | 0.14996
Mo, Molybdenum (ppm) | 2.3352 | 2.1998 | 2.4706 | 2.2125 | 2.4579
Na, Sodium (wt.%) | 1.9197 | 1.863 | 1.9763 | 1.8706 | 1.9688
Nb, Niobium (ppm) | 3.5365 | 3.3145 | 3.7584 | 3.3837 | 3.6892
Nd, Neodymium (ppm) | 8.2715 | 7.8109 | 8.7321 | 8.0209 | 8.522
Ni, Nickel (ppm) | 97.3957 | 94.4639 | 100.3275 | 95.3069 | 99.4846
P, Phosphorus (wt.%) | 0.042905 | 0.04162 | 0.04419 | 0.041939 | 0.043871
Pb, Lead (ppm) | 23.9222 | 22.8849 | 24.9594 | 23.3072 | 24.5372
Pr, Praseodymium (ppm) | 1.7099 | 1.6077 | 1.8121 | 1.6326 | 1.7873
Rb, Rubidium (ppm) | 14.0247 | 13.3494 | 14.7 | 13.5812 | 14.4683
Re, Rhenium (ppm) | 0.0031 | 0.0022057 | 0.0039943 | IND | IND
S, Sulphur (wt.%) | 0.37327 | 0.3629 | 0.38363 | 0.36359 | 0.38294
Sb, Antimony (ppm) | 1.3066 | 1.243 | 1.3703 | 1.2447 | 1.3685
Sc, Scandium (ppm) | 37.3814 | 35.8725 | 38.8903 | 36.4804 | 38.2824
Sm, Samarium (ppm) | 2.4659 | 2.2647 | 2.6671 | 2.3657 | 2.5661
Sn, Tin (ppm) | 1.1287 | 1.0039 | 1.2535 | IND | IND
Sr, Strontium (ppm) | 197.5049 | 191.7008 | 203.309 | 194.1028 | 200.907
Ta, Tantalum (ppm) | 0.2349 | 0.20613 | 0.26367 | 0.21559 | 0.2542
Tb, Terbium (ppm) | 0.53364 | 0.49366 | 0.57363 | 0.50948 | 0.55781
Te, Tellurium (ppm) | 0.095987 | 0.070836 | 0.12114 | IND | IND
Th, Thorium (ppm) | 0.92626 | 0.88115 | 0.97137 | 0.89222 | 0.96031
Ti, Titanium (wt.%) | 0.59663 | 0.58073 | 0.61253 | 0.58034 | 0.61292
Tl, Thallium (ppm) | 0.20919 | 0.19826 | 0.22013 | 0.19925 | 0.21913
Tm, Thulium (ppm) | 0.31632 | 0.29182 | 0.34082 | 0.2916 | 0.34105
U, Uranium (ppm) | 0.29707 | 0.28604 | 0.30809 | 0.28143 | 0.31271
V, Vanadium (ppm) | 262.0087 | 253.6578 | 270.3597 | 254.3062 | 269.7113
W, Tungsten (ppm) | 29.8841 | 28.1247 | 31.6436 | 28.8922 | 30.8761
Y, Yttrium (ppm) | 19.709 | 18.7653 | 20.6526 | 19.1841 | 20.2338
Yb, Ytterbium (ppm) | 2.0555 | 1.8929 | 2.2182 | 1.9618 | 2.1493
Zn, Zinc (ppm) | 137.7769 | 132.613 | 142.9408 | 134.7686 | 140.7852
Zr, Zirconium (ppm) | 50.1978 | 46.4712 | 53.9244 | 48.0744 | 52.3212
Aqua Regia Digestion
Ag, Silver (ppm) | 0.30617 | 0.28562 | 0.32672 | 0.29035 | 0.32199
Al, Aluminium (wt.%) | 3.188 | 3.0575 | 3.3184 | 3.1166 | 3.2594
As, Arsenic (ppm) | 58.9956 | 57.2808 | 60.7103 | 57.5693 | 60.4218
B, Boron (ppm) | 23.6901 | 18.4817 | 28.8986 | 22.0187 | 25.3616
Ba, Barium (ppm) | 39.1255 | 37.5319 | 40.7191 | 37.8305 | 40.4205
Be, Beryllium (ppm) | 0.20267 | 0.177 | 0.22833 | 0.18707 | 0.21827
Bi, Bismuth (ppm) | 0.041038 | 0.03418 | 0.047897 | IND | IND
Ca, Calcium (wt.%) | 2.2485 | 2.1156 | 2.3815 | 2.1862 | 2.3109
Cd, Cadmium (ppm) | 0.54255 | 0.51684 | 0.56826 | 0.51871 | 0.56639
Ce, Cerium (ppm) | 9.2757 | 8.9111 | 9.6404 | 8.9964 | 9.5551
Co, Cobalt (ppm) | 31.5941 | 30.5049 | 32.6833 | 30.9642 | 32.224
Cr, Chromium (ppm) | 28.3667 | 27.0637 | 29.6697 | 27.572 | 29.1614
Cs, Caesium (ppm) | 0.65227 | 0.62002 | 0.68452 | 0.63126 | 0.67328
Cu, Copper (ppm) | 160.2923 | 156.3425 | 164.242 | 157.0152 | 163.5694
Dy, Dysprosium (ppm) | 2.0555 | 1.7146 | 2.3964 | 1.9586 | 2.1524
Er, Erbium (ppm) | 1.19 | 0.98067 | 1.3994 | 1.1355 | 1.2446
Eu, Europium (ppm) | 0.46473 | 0.38759 | 0.54187 | 0.43513 | 0.49433
Fe, Iron (wt.%) | 5.7387 | 5.6215 | 5.8559 | 5.6172 | 5.8602
Ga, Gallium (ppm) | 9.9931 | 9.527 | 10.4592 | 9.7514 | 10.2347
Gd, Gadolinium (ppm) | 1.7984 | 1.5563 | 2.0404 | 1.7347 | 1.862
Ge, Germanium (ppm) | 0.11367 | 0.088982 | 0.13835 | IND | IND
Hf, Hafnium (ppm) | 0.41292 | 0.37334 | 0.45249 | 0.38747 | 0.43837
Ho, Holmium (ppm) | 0.42261 | 0.33639 | 0.50883 | 0.39546 | 0.44976
In, Indium (ppm) | 0.030076 | 0.026324 | 0.033827 | 0.021752 | 0.0384
K, Potassium (wt.%) | 0.14303 | 0.1372 | 0.14885 | 0.13897 | 0.14709
La, Lanthanum (ppm) | 4.0936 | 3.9422 | 4.245 | 3.9838 | 4.2034
Li, Lithium (ppm) | 10.4844 | 9.9422 | 11.0267 | 10.1868 | 10.7821
Lu, Lutetium (ppm) | 0.12723 | 0.10457 | 0.1499 | IND | IND
Mg, Magnesium (wt.%) | 1.8113 | 1.7567 | 1.8659 | 1.771 | 1.8516
Mn, Manganese (wt.%) | 0.080881 | 0.078049 | 0.083713 | 0.079425 | 0.082337
Mo, Molybdenum (ppm) | 2.2599 | 2.1308 | 2.3889 | 2.1753 | 2.3445
Na, Sodium (wt.%) | 0.22117 | 0.21074 | 0.23161 | 0.21084 | 0.23151
Nd, Neodymium (ppm) | 5.6347 | 5.1153 | 6.154 | 5.4159 | 5.8534
Ni, Nickel (ppm) | 67.6337 | 65.6277 | 69.6398 | 66.1102 | 69.1573
P, Phosphorus (wt.%) | 0.0419 | 0.040634 | 0.043166 | 0.040843 | 0.042957
Pb, Lead (ppm) | 23.5553 | 22.7109 | 24.3997 | 23.0968 | 24.0138
Pr, Praseodymium (ppm) | 1.1905 | 1.0754 | 1.3056 | 1.1348 | 1.2461
Rb, Rubidium (ppm) | 5.8834 | 5.5896 | 6.1773 | 5.7358 | 6.0311
Re, Rhenium (ppm) | 0.0018889 | 0.00092805 | 0.0028497 | IND | IND
S, Sulphur (wt.%) | 0.37174 | 0.36053 | 0.38295 | 0.3612 | 0.38229
Sb, Antimony (ppm) | 0.78655 | 0.68183 | 0.89128 | 0.72188 | 0.85122
Sc, Scandium (ppm) | 4.7753 | 4.4852 | 5.0654 | 4.6151 | 4.9355
Sm, Samarium (ppm) | 1.5135 | 1.2549 | 1.7721 | 1.4281 | 1.5989
Sn, Tin (ppm) | 0.62857 | 0.58049 | 0.67665 | 0.5762 | 0.68094
Sr, Strontium (ppm) | 36.6157 | 33.2681 | 39.9632 | 34.9258 | 38.3055
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.30405 | 0.28869 | 0.31942 | 0.29208 | 0.31603
Te, Tellurium (ppm) | 0.089437 | 0.065906 | 0.11297 | IND | IND
Th, Thorium (ppm) | 0.7188 | 0.68328 | 0.75432 | 0.69606 | 0.74154
Ti, Titanium (wt.%) | 0.31347 | 0.28628 | 0.34066 | 0.30137 | 0.32556
Tl, Thallium (ppm) | 0.10918 | 0.10109 | 0.11726 | IND | IND
Tm, Thulium (ppm) | 0.15996 | 0.12577 | 0.19415 | IND | IND
U, Uranium (ppm) | 0.1896 | 0.17797 | 0.20123 | 0.17134 | 0.20786
V, Vanadium (ppm) | 130.6089 | 122.2857 | 138.9321 | 126.8288 | 134.389
W, Tungsten (ppm) | 21.4458 | 19.7003 | 23.1914 | 20.7923 | 22.0994
Y, Yttrium (ppm) | 10.6648 | 10.3568 | 10.9729 | 10.4451 | 10.8846
Yb, Ytterbium (ppm) | 0.98243 | 0.83662 | 1.1282 | 0.91588 | 1.049
Zn, Zinc (ppm) | 123.4355 | 119.7665 | 127.1045 | 120.9681 | 125.9028
Zr, Zirconium (ppm) | 13.7535 | 12.6617 | 14.8453 | 13.1082 | 14.3989
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 233b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 18.8023 | Pt | ppb | 10.7285
4-Acid Digestion
B | ppm | 36.815 | Hg | ppm | < 1 | Se | ppm | 0.92201
Aqua Regia Digestion
Hg | ppm | 0.044556 | Pd | ppb | 17.6179 | Si | wt.% | 0.071459
Ir | ppm | < 0.003 | Pt | ppb | 8.3325
Nb | ppm | 0.11493 | Se | ppm | 0.59182
Borate Fusion XRF
Al2O3 | wt.% | 12.915 | MgO | wt.% | 6.535 | S | wt.% | 0.30895
CaO | wt.% | 9.17 | MnO | wt.% | 0.179 | SiO2 | wt.% | 51.825
Fe2O3 | wt.% | 11.645 | Na2O | wt.% | 2.55 | TiO2 | wt.% | 1.0375
K2O | wt.% | 0.6275 | P2O5 | wt.% | 0.094
Thermogravimetry
LOI1000 | wt.% | 3.175
Infrared Combustion
C | wt.% | 0.13 | S | wt.% | 0.35
Laser Ablation ICP-MS
Ag | ppm | 0.3 | Hf | ppm | 1.875 | Sm | ppm | 2.465
As | ppm | 58.8 | Ho | ppm | 0.8 | Sn | ppm | 1
Ba | ppm | 265 | In | ppm | 0.05 | Sr | ppm | 191.5
Be | ppm | 0.6 | La | ppm | 5.375 | Ta | ppm | 0.2
Bi | ppm | 0.04 | Lu | ppm | 0.335 | Tb | ppm | 0.545
Cd | ppm | 0.6 | Mn | wt.% | 0.157 | Te | ppm | < 0.2
Ce | ppm | 12.25 | Mo | ppm | 2.1 | Th | ppm | 0.99
Co | ppm | 44.75 | Nb | ppm | 3.5 | Ti | wt.% | 0.622
Cr | ppm | 177.5 | Nd | ppm | 8.545 | Tl | ppm | < 0.2
Cs | ppm | 0.88 | Ni | ppm | 105 | Tm | ppm | 0.35
Cu | ppm | 166 | Pb | ppm | 24 | U | ppm | 0.28
Dy | ppm | 3.66 | Pr | ppm | 1.805 | V | ppm | 274.5
Er | ppm | 2.255 | Rb | ppm | 13.6 | W | ppm | 30
Eu | ppm | 0.93 | Re | ppm | 0.025 | Y | ppm | 21.2
Ga | ppm | 15.2 | Sb | ppm | 1.3 | Yb | ppm | 2.245
Gd | ppm | 3.085 | Sc | ppm | 38.85 | Zn | ppm | 145
Ge | ppm | 1.425 | Se | ppm | < 5 | Zr | ppm | 65
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 233b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 1.075 | 0.042686 | 0.98962 | 1.1604 | 0.94694 | 1.2031 | 3.971% | 7.942% | 11.91% | 1.0212 | 1.1287
PhotonAssay
Au, ppm | 1.0698 | 0.042681 | 0.98448 | 1.1552 | 0.9418 | 1.1979 | 3.989% | 7.979% | 11.97% | 1.0164 | 1.1233
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.0124 | 0.037344 | 0.93767 | 1.087 | 0.90033 | 1.1244 | 3.689% | 7.378% | 11.07% | 0.96174 | 1.063
Cyanide Leach
Au, ppm | 1.0053 | 0.055152 | 0.89505 | 1.1157 | 0.8399 | 1.1708 | 5.486% | 10.97% | 16.46% | 0.95508 | 1.0556
4-Acid Digestion
Ag, ppm | 0.31065 | 0.015924 | 0.2788 | 0.3425 | 0.26287 | 0.35842 | 5.126% | 10.25% | 15.38% | 0.29511 | 0.32618
Al, wt.% | 6.7985 | 0.26383 | 6.2709 | 7.3262 | 6.007 | 7.59 | 3.881% | 7.761% | 11.64% | 6.4586 | 7.1385
As, ppm | 59.856 | 2.4841 | 54.8878 | 64.8241 | 52.4037 | 67.3082 | 4.15% | 8.3% | 12.45% | 56.8632 | 62.8488
Ba, ppm | 270.0018 | 10.8021 | 248.3976 | 291.606 | 237.5955 | 302.4081 | 4.001% | 8.002% | 12% | 256.5017 | 283.5018
Be, ppm | 0.39138 | 0.045061 | 0.30126 | 0.4815 | 0.2562 | 0.52657 | 11.51% | 23.03% | 34.54% | 0.37181 | 0.41095
Bi, ppm | 0.045568 | 0.0096711 | 0.026226 | 0.064911 | 0.016555 | 0.074582 | 21.22% | 42.45% | 63.67% | 0.04329 | 0.047847
Ca, wt.% | 6.3251 | 0.24551 | 5.8341 | 6.8161 | 5.5886 | 7.0616 | 3.882% | 7.763% | 11.64% | 6.0088 | 6.6413
Cd, ppm | 0.58174 | 0.049826 | 0.48209 | 0.6814 | 0.43227 | 0.73122 | 8.565% | 17.13% | 25.69% | 0.55266 | 0.61083
Ce, ppm | 12.5654 | 0.72497 | 11.1155 | 14.0154 | 10.3905 | 14.7404 | 5.77% | 11.54% | 17.31% | 11.9372 | 13.1937
Co, ppm | 43.5587 | 1.6864 | 40.186 | 46.9314 | 38.4996 | 48.6178 | 3.871% | 7.743% | 11.61% | 41.3808 | 45.7366
Cr, ppm | 146.0186 | 14.7075 | 116.6036 | 175.4335 | 101.8962 | 190.141 | 10.07% | 20.14% | 30.22% | 138.7176 | 153.3195
Cs, ppm | 0.89133 | 0.057589 | 0.77615 | 1.0065 | 0.71856 | 1.0641 | 6.461% | 12.92% | 19.38% | 0.84676 | 0.9359
Cu, ppm | 160.8734 | 6.756 | 147.3614 | 174.3853 | 140.6055 | 181.1413 | 4.2% | 8.399% | 12.6% | 152.8297 | 168.9171
Dy, ppm | 3.6044 | 0.15065 | 3.3031 | 3.9057 | 3.1524 | 4.0563 | 4.18% | 8.359% | 12.54% | 3.4242 | 3.7846
Er, ppm | 2.1847 | 0.12928 | 1.9262 | 2.4433 | 1.7969 | 2.5726 | 5.917% | 11.83% | 17.75% | 2.0755 | 2.294
Eu, ppm | 0.89567 | 0.052918 | 0.78984 | 1.0015 | 0.73692 | 1.0544 | 5.908% | 11.82% | 17.72% | 0.85089 | 0.94046
Fe, wt.% | 7.8431 | 0.33341 | 7.1762 | 8.5099 | 6.8428 | 8.8433 | 4.251% | 8.502% | 12.75% | 7.4509 | 8.2352
Ga, ppm | 15.0173 | 1.1116 | 12.794 | 17.2405 | 11.6824 | 18.3522 | 7.402% | 14.8% | 22.21% | 14.2664 | 15.7681
Gd, ppm | 2.9573 | 0.23835 | 2.4806 | 3.434 | 2.2422 | 3.6723 | 8.06% | 16.12% | 24.18% | 2.8094 | 3.1051
Ge, ppm | 0.08275 | 0.017471 | 0.047807 | 0.11769 | 0.030336 | 0.13516 | 21.11% | 42.23% | 63.34% | 0.078613 | 0.086888
Hf, ppm | 1.5976 | 0.20278 | 1.192 | 2.0032 | 0.98926 | 2.206 | 12.69% | 25.39% | 38.08% | 1.5177 | 1.6775
Ho, ppm | 0.74299 | 0.044229 | 0.65454 | 0.83145 | 0.61031 | 0.87568 | 5.953% | 11.91% | 17.86% | 0.70584 | 0.78014
In, ppm | 0.073004 | 0.0085121 | 0.05598 | 0.090029 | 0.047468 | 0.098541 | 11.66% | 23.32% | 34.98% | 0.069354 | 0.076655
K, wt.% | 0.54135 | 0.026179 | 0.489 | 0.59371 | 0.46282 | 0.61989 | 4.836% | 9.672% | 14.51% | 0.51429 | 0.56842
La, ppm | 5.301 | 0.22494 | 4.8511 | 5.7509 | 4.6262 | 5.9759 | 4.243% | 8.487% | 12.73% | 5.036 | 5.5661
Li, ppm | 12.192 | 1.2761 | 9.6398 | 14.7441 | 8.3637 | 16.0202 | 10.47% | 20.93% | 31.4% | 11.5824 | 12.8016
Lu, ppm | 0.32391 | 0.019964 | 0.28399 | 0.36384 | 0.26402 | 0.38381 | 6.163% | 12.33% | 18.49% | 0.30772 | 0.34011
Mg, wt.% | 3.8702 | 0.14435 | 3.5815 | 4.1589 | 3.4372 | 4.3033 | 3.73% | 7.46% | 11.19% | 3.6767 | 4.0637
Mn, wt.% | 0.14715 | 0.0051453 | 0.13686 | 0.15744 | 0.13172 | 0.16259 | 3.497% | 6.993% | 10.49% | 0.13979 | 0.15451
Mo, ppm | 2.3352 | 0.19706 | 1.9411 | 2.7293 | 1.744 | 2.9264 | 8.438% | 16.88% | 25.32% | 2.2184 | 2.452
Na, wt.% | 1.9197 | 0.071467 | 1.7767 | 2.0626 | 1.7053 | 2.1341 | 3.723% | 7.446% | 11.17% | 1.8237 | 2.0157
Nb, ppm | 3.5365 | 0.32403 | 2.8884 | 4.1845 | 2.5644 | 4.5086 | 9.162% | 18.32% | 27.49% | 3.3597 | 3.7133
Nd, ppm | 8.2715 | 0.36327 | 7.545 | 8.998 | 7.1817 | 9.3613 | 4.392% | 8.784% | 13.18% | 7.8579 | 8.6851
Ni, ppm | 97.3957 | 4.6433 | 88.1091 | 106.6824 | 83.4657 | 111.3257 | 4.767% | 9.535% | 14.3% | 92.5259 | 102.2655
P, wt.% | 0.042905 | 0.0014709 | 0.039963 | 0.045847 | 0.038492 | 0.047318 | 3.428% | 6.856% | 10.28% | 0.04076 | 0.04505
Pb, ppm | 23.9222 | 1.7445 | 20.4331 | 27.4112 | 18.6886 | 29.1557 | 7.293% | 14.59% | 21.88% | 22.7261 | 25.1183
Pr, ppm | 1.7099 | 0.087614 | 1.5347 | 1.8851 | 1.4471 | 1.9728 | 5.124% | 10.25% | 15.37% | 1.6244 | 1.7954
Rb, ppm | 14.0247 | 0.69568 | 12.6333 | 15.4161 | 11.9377 | 16.1118 | 4.96% | 9.921% | 14.88% | 13.3235 | 14.726
Re, ppm | 0.0031 | 0.00091306 | 0.0012739 | 0.0049261 | 0.00036081 | 0.0058392 | 29.45% | 58.91% | 88.36% | 0.002945 | 0.003255
S, wt.% | 0.37327 | 0.013278 | 0.34671 | 0.39982 | 0.33343 | 0.4131 | 3.557% | 7.115% | 10.67% | 0.3546 | 0.39193
Sb, ppm | 1.3066 | 0.079831 | 1.147 | 1.4663 | 1.0671 | 1.5461 | 6.11% | 12.22% | 18.33% | 1.2413 | 1.372
Sc, ppm | 37.3814 | 2.3172 | 32.747 | 42.0158 | 30.4298 | 44.333 | 6.199% | 12.4% | 18.6% | 35.5123 | 39.2505
Sm, ppm | 2.4659 | 0.11866 | 2.2286 | 2.7032 | 2.1099 | 2.8219 | 4.812% | 9.625% | 14.44% | 2.3426 | 2.5892
Sn, ppm | 1.1287 | 0.18235 | 0.764 | 1.4934 | 0.58164 | 1.6758 | 16.16% | 32.31% | 48.47% | 1.0723 | 1.1851
Sr, ppm | 197.5049 | 8.2558 | 180.9933 | 214.0166 | 172.7374 | 222.2724 | 4.18% | 8.36% | 12.54% | 187.6297 | 207.3802
Ta, ppm | 0.2349 | 0.040321 | 0.15426 | 0.31554 | 0.11394 | 0.35586 | 17.17% | 34.33% | 51.5% | 0.22315 | 0.24664
Tb, ppm | 0.53364 | 0.038358 | 0.45693 | 0.61036 | 0.41857 | 0.64872 | 7.188% | 14.38% | 21.56% | 0.50696 | 0.56033
Te, ppm | 0.095987 | 0.017579 | 0.060829 | 0.13114 | 0.043251 | 0.14872 | 18.31% | 36.63% | 54.94% | 0.091187 | 0.10079
Th, ppm | 0.92626 | 0.058217 | 0.80983 | 1.0427 | 0.75161 | 1.1009 | 6.285% | 12.57% | 18.86% | 0.87995 | 0.97258
Ti, wt.% | 0.59663 | 0.010714 | 0.5752 | 0.61806 | 0.56449 | 0.62877 | 1.796% | 3.592% | 5.387% | 0.5668 | 0.62646
Tl, ppm | 0.20919 | 0.013419 | 0.18235 | 0.23603 | 0.16894 | 0.24945 | 6.415% | 12.83% | 19.24% | 0.19873 | 0.21965
Tm, ppm | 0.31632 | 0.021032 | 0.27426 | 0.35839 | 0.25323 | 0.37942 | 6.649% | 13.3% | 19.95% | 0.30051 | 0.33214
U, ppm | 0.29707 | 0.0085342 | 0.28 | 0.31414 | 0.27147 | 0.32267 | 2.873% | 5.746% | 8.618% | 0.28221 | 0.31192
V, ppm | 262.0087 | 10.0529 | 241.903 | 282.1145 | 231.8501 | 292.1674 | 3.837% | 7.674% | 11.51% | 248.9083 | 275.1092
W, ppm | 29.8841 | 2.6427 | 24.5988 | 35.1695 | 21.9561 | 37.8121 | 8.843% | 17.69% | 26.53% | 28.3899 | 31.3783
Y, ppm | 19.709 | 1.1245 | 17.46 | 21.9579 | 16.3356 | 23.0824 | 5.705% | 11.41% | 17.12% | 18.7235 | 20.6944
Yb, ppm | 2.0555 | 0.1732 | 1.7091 | 2.4019 | 1.5359 | 2.5752 | 8.426% | 16.85% | 25.28% | 1.9528 | 2.1583
Zn, ppm | 137.7769 | 7.8928 | 121.9913 | 153.5625 | 114.0985 | 161.4553 | 5.729% | 11.46% | 17.19% | 130.888 | 144.6657
Zr, ppm | 50.1978 | 8.2027 | 33.7925 | 66.6031 | 25.5899 | 74.8058 | 16.34% | 32.68% | 49.02% | 47.6879 | 52.7077
Aqua Regia Digestion
Ag, ppm | 0.30617 | 0.01874 | 0.26869 | 0.34365 | 0.24995 | 0.36239 | 6.121% | 12.24% | 18.36% | 0.29086 | 0.32148
Al, wt.% | 3.188 | 0.26022 | 2.6675 | 3.7084 | 2.4073 | 3.9686 | 8.163% | 16.33% | 24.49% | 3.0286 | 3.3474
As, ppm | 58.9956 | 2.5377 | 53.9201 | 64.071 | 51.3824 | 66.6087 | 4.302% | 8.603% | 12.9% | 56.0458 | 61.9453
B, ppm | 23.6901 | 7.2285 | 9.2331 | 38.1472 | 2.0045 | 45.3758 | 30.51% | 61.03% | 91.54% | 22.5056 | 24.8747
Ba, ppm | 39.1255 | 1.9681 | 35.1892 | 43.0618 | 33.2211 | 45.0299 | 5.03% | 10.06% | 15.09% | 37.1692 | 41.0818
Be, ppm | 0.20267 | 0.029289 | 0.14409 | 0.26125 | 0.1148 | 0.29054 | 14.45% | 28.9% | 43.36% | 0.19253 | 0.2128
Bi, ppm | 0.041038 | 0.0066008 | 0.027837 | 0.05424 | 0.021236 | 0.060841 | 16.08% | 32.17% | 48.25% | 0.038986 | 0.04309
Ca, wt.% | 2.2485 | 0.27674 | 1.6951 | 2.802 | 1.4183 | 3.0787 | 12.31% | 24.61% | 36.92% | 2.1361 | 2.361
Cd, ppm | 0.54255 | 0.036022 | 0.47051 | 0.61459 | 0.43448 | 0.65061 | 6.639% | 13.28% | 19.92% | 0.51542 | 0.56968
Ce, ppm | 9.2757 | 0.51404 | 8.2477 | 10.3038 | 7.7336 | 10.8179 | 5.542% | 11.08% | 16.63% | 8.812 | 9.7395
Co, ppm | 31.5941 | 1.8104 | 27.9733 | 35.2149 | 26.1629 | 37.0253 | 5.73% | 11.46% | 17.19% | 30.0144 | 33.1738
Cr, ppm | 28.3667 | 2.4858 | 23.395 | 33.3384 | 20.9092 | 35.8242 | 8.763% | 17.53% | 26.29% | 26.9484 | 29.785
Cs, ppm | 0.65227 | 0.043604 | 0.56506 | 0.73948 | 0.52146 | 0.78308 | 6.685% | 13.37% | 20.05% | 0.61966 | 0.68488
Cu, ppm | 160.2923 | 5.8611 | 148.5701 | 172.0144 | 142.709 | 177.8755 | 3.657% | 7.313% | 10.97% | 152.2776 | 168.3069
Dy, ppm | 2.0555 | 0.29352 | 1.4684 | 2.6425 | 1.1749 | 2.936 | 14.28% | 28.56% | 42.84% | 1.9527 | 2.1582
Er, ppm | 1.19 | 0.18301 | 0.82404 | 1.5561 | 0.64103 | 1.7391 | 15.38% | 30.76% | 46.13% | 1.1305 | 1.2496
Eu, ppm | 0.46473 | 0.057472 | 0.34978 | 0.57967 | 0.29231 | 0.63714 | 12.37% | 24.73% | 37.1% | 0.44149 | 0.48796
Fe, wt.% | 5.7387 | 0.18741 | 5.3639 | 6.1135 | 5.1765 | 6.301 | 3.266% | 6.532% | 9.797% | 5.4518 | 6.0257
Ga, ppm | 9.9931 | 0.78298 | 8.4271 | 11.559 | 7.6442 | 12.342 | 7.835% | 15.67% | 23.51% | 9.4934 | 10.4927
Gd, ppm | 1.7984 | 0.2052 | 1.388 | 2.2087 | 1.1828 | 2.4139 | 11.41% | 22.82% | 34.23% | 1.7084 | 1.8883
Ge, ppm | 0.11367 | 0.022697 | 0.068273 | 0.15906 | 0.045576 | 0.18176 | 19.97% | 39.94% | 59.9% | 0.10798 | 0.11935
Hf, ppm | 0.41292 | 0.044117 | 0.32468 | 0.50115 | 0.28056 | 0.54527 | 10.68% | 21.37% | 32.05% | 0.39227 | 0.43356
Ho, ppm | 0.42261 | 0.068097 | 0.28642 | 0.55881 | 0.21832 | 0.6269 | 16.11% | 32.23% | 48.34% | 0.40148 | 0.44374
In, ppm | 0.030076 | 0.0034513 | 0.023173 | 0.036978 | 0.019722 | 0.04043 | 11.48% | 22.95% | 34.43% | 0.028572 | 0.03158
K, wt.% | 0.14303 | 0.0056506 | 0.13173 | 0.15433 | 0.12607 | 0.15998 | 3.951% | 7.901% | 11.85% | 0.13587 | 0.15018
La, ppm | 4.0936 | 0.17739 | 3.7388 | 4.4484 | 3.5614 | 4.6258 | 4.333% | 8.667% | 13% | 3.8889 | 4.2983
Li, ppm | 10.4844 | 0.79788 | 8.8887 | 12.0802 | 8.0908 | 12.8781 | 7.61% | 15.22% | 22.83% | 9.9602 | 11.0087
Lu, ppm | 0.12723 | 0.014203 | 0.098826 | 0.15564 | 0.084623 | 0.16984 | 11.16% | 22.33% | 33.49% | 0.12087 | 0.13359
Mg, wt.% | 1.8113 | 0.094005 | 1.6233 | 1.9993 | 1.5293 | 2.0933 | 5.19% | 10.38% | 15.57% | 1.7207 | 1.9019
Mn, wt.% | 0.080881 | 0.0059323 | 0.069016 | 0.092746 | 0.063084 | 0.098678 | 7.335% | 14.67% | 22% | 0.076837 | 0.084925
Mo, ppm | 2.2599 | 0.1199 | 2.0201 | 2.4997 | 1.9002 | 2.6196 | 5.306% | 10.61% | 15.92% | 2.1469 | 2.3729
Na, wt.% | 0.22117 | 0.015278 | 0.19062 | 0.25173 | 0.17534 | 0.26701 | 6.908% | 13.82% | 20.72% | 0.21012 | 0.23223
Nd, ppm | 5.6347 | 0.4391 | 4.7565 | 6.5129 | 4.3174 | 6.9519 | 7.793% | 15.59% | 23.38% | 5.3529 | 5.9164
Ni, ppm | 67.6337 | 3.13 | 61.3737 | 73.8938 | 58.2436 | 77.0238 | 4.628% | 9.256% | 13.88% | 64.2521 | 71.0154
P, wt.% | 0.0419 | 0.0018382 | 0.038223 | 0.045576 | 0.036385 | 0.047414 | 4.387% | 8.774% | 13.16% | 0.039805 | 0.043995
Pb, ppm | 23.5553 | 1.1406 | 21.2741 | 25.8365 | 20.1335 | 26.9771 | 4.842% | 9.684% | 14.53% | 22.3775 | 24.7331
Pr, ppm | 1.1905 | 0.093996 | 1.0025 | 1.3785 | 0.9085 | 1.4725 | 7.896% | 15.79% | 23.69% | 1.131 | 1.25
Rb, ppm | 5.8834 | 0.44907 | 4.9853 | 6.7816 | 4.5362 | 7.2307 | 7.633% | 15.27% | 22.9% | 5.5893 | 6.1776
Re, ppm | 0.0018889 | 0.00050157 | 0.00088575 | 0.002892 | 0.00038418 | 0.0033936 | 26.55% | 53.11% | 79.66% | 0.0017944 | 0.0019833
S, wt.% | 0.37174 | 0.014239 | 0.34326 | 0.40022 | 0.32902 | 0.41446 | 3.83% | 7.661% | 11.49% | 0.35316 | 0.39033
Sb, ppm | 0.78655 | 0.19995 | 0.38665 | 1.1865 | 0.1867 | 1.3864 | 25.42% | 50.84% | 76.26% | 0.74723 | 0.82588
Sc, ppm | 4.7753 | 0.40929 | 3.9567 | 5.5939 | 3.5475 | 6.0032 | 8.571% | 17.14% | 25.71% | 4.5365 | 5.0141
Sm, ppm | 1.5135 | 0.16312 | 1.1872 | 1.8397 | 1.0241 | 2.0028 | 10.78% | 21.56% | 32.33% | 1.4378 | 1.5892
Sn, ppm | 0.62857 | 0.051206 | 0.52616 | 0.73098 | 0.47495 | 0.78219 | 8.146% | 16.29% | 24.44% | 0.59714 | 0.66
Sr, ppm | 36.6157 | 6.6576 | 23.3004 | 49.9309 | 16.6428 | 56.5885 | 18.18% | 36.36% | 54.55% | 34.7849 | 38.4464
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.30405 | 0.010704 | 0.28265 | 0.32546 | 0.27194 | 0.33617 | 3.52% | 7.041% | 10.56% | 0.28885 | 0.31926
Te, ppm | 0.089437 | 0.014752 | 0.059933 | 0.11894 | 0.045181 | 0.13369 | 16.49% | 32.99% | 49.48% | 0.084965 | 0.093909
Th, ppm | 0.7188 | 0.070841 | 0.57712 | 0.86049 | 0.50628 | 0.93133 | 9.855% | 19.71% | 29.57% | 0.68286 | 0.75474
Ti, wt.% | 0.31347 | 0.057485 | 0.1985 | 0.42844 | 0.14101 | 0.48592 | 18.34% | 36.68% | 55.02% | 0.29779 | 0.32914
Tl, ppm | 0.10918 | 0.0063672 | 0.096441 | 0.12191 | 0.090074 | 0.12828 | 5.832% | 11.66% | 17.5% | 0.10372 | 0.11463
Tm, ppm | 0.15996 | 0.02747 | 0.10502 | 0.2149 | 0.077551 | 0.24237 | 17.17% | 34.35% | 51.52% | 0.15196 | 0.16796
U, ppm | 0.1896 | 0.012754 | 0.16409 | 0.21511 | 0.15134 | 0.22786 | 6.727% | 13.45% | 20.18% | 0.18012 | 0.19908
V, ppm | 130.6089 | 15.585 | 99.439 | 161.7788 | 83.854 | 177.3638 | 11.93% | 23.87% | 35.8% | 124.0785 | 137.1394
W, ppm | 21.4458 | 3.5405 | 14.3648 | 28.5269 | 10.8243 | 32.0674 | 16.51% | 33.02% | 49.53% | 20.3736 | 22.5181
Y, ppm | 10.6648 | 0.43155 | 9.8017 | 11.5279 | 9.3702 | 11.9595 | 4.046% | 8.093% | 12.14% | 10.1316 | 11.1981
Yb, ppm | 0.98243 | 0.14929 | 0.68384 | 1.281 | 0.53455 | 1.4303 | 15.2% | 30.39% | 45.59% | 0.9333 | 1.0315
Zn, ppm | 123.4355 | 7.1655 | 109.1046 | 137.7664 | 101.9391 | 144.9318 | 5.805% | 11.61% | 17.42% | 117.2637 | 129.6072
Zr, ppm | 13.7535 | 1.8678 | 10.018 | 17.489 | 8.1502 | 19.3568 | 13.58% | 27.16% | 40.74% | 13.0658 | 14.4412
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 233b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
N.A. | Not Applicable (Lab 1.27)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 233b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
Nagrom, Perth, WA, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 233b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 17 | 17 | 238537 | 0.084515 | 1.0923 | 1.1036
2 | 1 | ANSLu | 1 | 11 | 11 | 238538 | 0.083356 | 1.0687 | 1.1023
2 | 1 | ANSLu | 1 | 4 | 4 | 238539 | 0.087439 | 1.0844 | 1.1031
2 | 1 | ANSLu | 1 | 13 | 13 | 238540 | 0.084085 | 1.0627 | 1.102
2 | 1 | ANSLu | 1 | 14 | 14 | 238541 | 0.086652 | 1.0707 | 1.1024
2 | 1 | ANSLu | 1 | 20 | 20 | 238542 | 0.086169 | 1.0665 | 1.1022
2 | 1 | ANSLu | 1 | 1 | 1 | 238543 | 0.087196 | 1.077 | 1.1027
2 | 1 | ANSLu | 1 | 10 | 10 | 238544 | 0.0834 | 1.1935 | 1.109
2 | 1 | ANSLu | 1 | 19 | 19 | 238545 | 0.088102 | 1.2162 | 1.1102
2 | 1 | ANSLu | 1 | 3 | 3 | 238546 | 0.083687 | 1.0784 | 1.1028
2 | 1 | ANSLu | 1 | 7 | 7 | 238547 | 0.087694 | 1.143 | 1.1063
2 | 1 | ANSLu | 1 | 18 | 18 | 238548 | 0.087088 | 1.1099 | 1.1045
2 | 1 | ANSLu | 1 | 8 | 8 | 238549 | 0.087428 | 1.0642 | 1.1021
2 | 1 | ANSLu | 1 | 9 | 9 | 238550 | 0.086067 | 1.0695 | 1.1023
2 | 1 | ANSLu | 1 | 15 | 15 | 238551 | 0.087137 | 1.1245 | 1.1053
2 | 1 | ANSLu | 1 | 6 | 6 | 238552 | 0.086993 | 1.0808 | 1.1029
2 | 1 | ANSLu | 1 | 12 | 12 | 238553 | 0.085587 | 1.082 | 1.103
2 | 1 | ANSLu | 1 | 16 | 16 | 238554 | 0.08617 | 1.0924 | 1.1036
2 | 1 | ANSLu | 1 | 5 | 5 | 238555 | 0.084109 | 1.1418 | 1.1062
2 | 1 | ANSLu | 1 | 2 | 2 | 238556 | 0.083261 | 1.1651 | 1.1074
Mean | [6 empty] | 0.085807 | 1.1042 | 1.1042
Median | [7 empty] | 1.0832 | 1.1031
SD | [7 empty] | 0.04533 | 0.0024243
RSD | [7 empty] | 4.105% | 0.2196%
RSD Upscale Factor | [7 empty] | 0.053481
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 233b (Certified Value 1.07 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 30g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 50g | 50g | 25g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 15g | 30g | 30g | 25g | 30g | 50g | 30g | 30g | 30g | 3
| 1 | 1 | 1.077 | 0.967 | 1.066 | 1.078 | 1.1125 | 1.1304 | 1.02 | 1.05 | 1.08 | 1.02 | 1.12 | 1.073 | 1.14 | 0.944 | 1.08 | 1.055 | 1.06 | 1.038 | 1.03 | 0.946 | 1.04 | 1.057 | 0.944 | 1.155 | 1.0994 | 1.045 | 1.1338 | 1.089 | 1.07 | 1.094 | 1.062 | 1.17 | 1
| 1 | 2 | 1.1651 | 0.922 | 1.121 | 1.072 | 1.0737 | 1.084 | 1.02 | 1.07 | 1.07 | 1.06 | 1.13 | 1.069 | 1.12 | 1.038 | 1.065 | 1.08 | 1.08 | 1.073 | 1.065 | 1.012 | 1.05 | 1.066 | 1.04 | 1.064 | 1.1025 | 1.033 | 1.1114 | 1.099 | 1.06 | 1.053 | 1.067 | 1.18 | 0
| 1 | 3 | 1.0784 | 1.089 | 1.098 | 1.102 | 1.1134 | 1.1541 | 1.04 | 1.07 | 1.11 | 1.01 | 1.14 | 1.061 | 1.142 | 0.993 | 1.07 | 1.09 | 1.05 | 1.093 | 1.04 | 1.017 | 1.06 | 1.039 | 1.04 | 1.069 | 1.1015 | 1.074 | 1.12 | 1.096 | 1.04 | 1.062 | 1.051 | 1.17 | 16
| 1 | 4 | 1.0844 | 0.903 | 1.107 | 1.102 | 1.1212 | 1.1177 | 1.04 | 1.08 | 1.05 | 1.05 | 1.11 | 1.059 | 1.12 | 0.982 | 1.06 | 1.095 | 1.085 | 1.075 | 1.015 | 1.044 | 1.04 | 1.046 | 1.03 | 1.126 | 1.1025 | 1.061 | 1.1075 | 1.09 | 1.02 | 1.083 | 1.049 | 1.16 | 1.075
| 1 | 5 | 1.1418 | 1.084 | 1.088 | 1.066 | 1.1096 | 1.1308 | 1.02 | 1.08 | 1.07 | 1 | 1.12 | 1.086 | 1.154 | 1.022 | 1.07 | 1.085 | 1.075 | 1.097 | 1.04 | 1.053 | 1.05 | 1.058 | 1.04 | 1.168 | 1.0941 | 1.07 | 1.1096 | 1.096 | 1.05 | 1.078 | 1.087 | 1.17 | 7
| 1 | 6 | 1.0808 | 1.076 | 1.092 | 1.122 | 1.1057 | 1.1331 | 1.02 | 1.06 | 1.08 | 1.01 | 1.14 | 1.075 | 1.12 | 1 | 1.085 | 1.075 | 1.09 | 1.104 | 1 | 0.966 | 1.06 | 1.115 | 1.04 | 1.099 | 1.0973 | 1.056 | 1.1122 | 1.09 | 1.07 | 1.073 | 1.07 | 1.16
|  | 7 | 2.00: 1.143
|  | 8 | 2.00: 1.0642
|  | 9 | 2.00: 1.0695
|  | 10 | 2.00: 1.1935
|  | 11 | 2.00: 1.0687
|  | 12 | 2.00: 1.082
|  | 13 | 2.00: 1.0627
|  | 14 | 2.00: 1.0707
|  | 15 | 2.00: 1.1245
|  | 16 | 2.00: 1.0924
|  | 17 | 2.00: 1.0923
|  | 18 | 2.00: 1.1099
|  | 19 | 2.00: 1.2162
|  | 20 | 2.00: 1.0665
| Mean |  | 1.1042 | 1.0068 | 1.0953 | 1.0903 | 1.106 | 1.125 | 1.0267 | 1.0683 | 1.0767 | 1.025 | 1.1267 | 1.0705 | 1.1327 | 0.9965 | 1.0717 | 1.08 | 1.0733 | 1.08 | 1.0317 | 1.0063 | 1.05 | 1.0635 | 1.0223 | 1.1135 | 1.0995 | 1.0565 | 1.1157 | 1.0933 | 1.0517 | 1.0738 | 1.0643 | 1.1683
| Median |  | 1.0832 | 1.0215 | 1.095 | 1.09 | 1.111 | 1.1306 | 1.02 | 1.07 | 1.075 | 1.015 | 1.125 | 1.071 | 1.13 | 0.9965 | 1.07 | 1.0825 | 1.0775 | 1.084 | 1.035 | 1.0145 | 1.05 | 1.0575 | 1.04 | 1.1125 | 1.1004 | 1.0585 | 1.1118 | 1.093 | 1.055 | 1.0755 | 1.0645 | 1.17
| Std Dev. |  | 0.04533 | 0.086087 | 0.018587 | 0.021704 | 0.016651 | 0.023257 | 0.010328 | 0.01169 | 0.019664 | 0.02429 | 0.012111 | 0.0099146 | 0.014679 | 0.032752 | 0.0093095 | 0.014142 | 0.015384 | 0.023967 | 0.022509 | 0.042448 | 0.0089443 | 0.026972 | 0.038583 | 0.043565 | 0.0033479 | 0.01545 | 0.0097968 | 0.0041793 | 0.019408 | 0.014716 | 0.013938 | 0.0075277
| Rel.Std.Dev. |  | 4.105% | 8.55% | 1.697% | 1.991% | 1.506% | 2.067% | 1.006% | 1.094% | 1.826% | 2.37% | 1.075% | 0.9262% | 1.296% | 3.287% | 0.8687% | 1.309% | 1.433% | 2.219% | 2.182% | 4.218% | 0.8518% | 2.536% | 3.774% | 3.912% | 0.3045% | 1.462% | 0.8781% | 0.3823% | 1.845% | 1.37% | 1.31% | 0.6443%
| PDM3 |  | 2.715% | -6.341% | 1.892% | 1.427% | 2.884% | 4.652% | -4.496% | -0.6199% | 0.1553% | -4.651% | 4.806% | -0.4184% | 5.365% | -7.302% | -0.3099% | 0.4653% | -0.1548% | 0.4653% | -4.031% | -6.387% | -2.325% | -1.07% | -4.899% | 3.582% | 2.282% | -1.721% | 3.789% | 1.706% | -2.17% | -0.1083% | -0.992% | 8.682%
| Z-Score (Absolute) |  | NA | 1.92 | 0.64 | 0.49 | 0.94 | 1.49 | 1.35 | 0.14 | 0.1 | 1.4 | 1.54 | 0.08 | 1.71 | 2.22 | 0.05 | 0.19 | 0 | 0.19 | 1.2 | 1.94 | 0.67 | 0.28 | 1.47 | 1.16 | 0.76 | 0.49 | 1.23 | 0.58 | 0.63 | 0.01 | 0.26 | 2.75
| Analytical results for Pd in OREAS 233b (Indicative Value 18.8 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.28 | ppb
|  |  | 2.00: FA*MS | 1
|  |  | 2.00: 25g | 1
| 1 | 1 | 2.00: 18.254 | 1
| 1 | 2 | 2.00: 19.0969 | 1
| 1 | 3 | 2.00: 20.6768 | 16
| 1 | 4 | 2.00: 18.0327 | 18.8023
| 1 | 5 | 2.00: 19.0229 | 7
| 1 | 6 | 2.00: 17.7304
| Mean |  | 2.00: 18.8023
| Median |  | 2.00: 18.6384
| Std Dev. |  | 2.00: 1.0669
| Rel.Std.Dev. |  | 2.00: 5.674%
| PDM3 |  | 2.00: -1.665e-13%
Z-Score (Absolute) | bar: 2.00
| Analytical results for Pt in OREAS 233b (Indicative Value 10.7 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.28 | ppb
|  |  | 2.00: FA*MS | 1
|  |  | 2.00: 25g | 1
| 1 | 1 | 2.00: 10.6825 | 1
| 1 | 2 | 2.00: 10.794 | 1
| 1 | 3 | 2.00: 10.7582 | 16
| 1 | 4 | 2.00: 10.8634 | 10.7285
| 1 | 5 | 2.00: 10.4946 | 7
| 1 | 6 | 2.00: 10.7783
| Mean |  | 2.00: 10.7285
| Median |  | 2.00: 10.7682
| Std Dev. |  | 2.00: 0.1286
| Rel.Std.Dev. |  | 2.00: 1.199%
| PDM3 |  | 2.00: 2.887e-13%
Z-Score (Absolute) | bar: 2.00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 233b (Certified Value 1.07 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | 5.05 | 5.06 | 5.07 | 6.01 | 6.02 | 6.03 | 6.04 | ppm
|  |  | INAA | Raw*PA | Various [a] | Raw*PA | Various [b] | Various [c] | Raw*PA | Various [d] | Raw*PA | Raw*PA | Raw*PA | Various [e] | 2
|  |  | 0.085g | 400g | see footer | 400g | see footer | see footer | 400g | see footer | 450g | 400g | 450g | see footer | 3
| 1 | 1 | 1.077 | 1.12 | 1.13 |  | 1.0488 | 1.047 | 1.0607 | 1.023 | 1.106 | 1.135 | 1 | 1.061 | 1
| 1 | 2 | 1.1651 | 1.04 | 1.05 |  | 1.0139 | 1.066 | 1.0734 | 1.003 | 1.058 | 1.066 | 1.093 | 1.1 | 0
| 1 | 3 | 1.0784 | 1.1 | 1.04 |  | 1.0805 | 1.043 | 1.041 | 1.08 | 1.122 | 1.132 | 1.099 | 1.117 | 16
| 1 | 4 | 1.0844 | 1.08 | 1.08 |  | 0.98461 | 1.074 | 1.0011 | 1.031 | 1.037 | 1.101 | 1.037 | 1.057 | 1.0698
| 1 | 5 | 1.1418 | 1.05 | 1.11 |  | 1.0899 | 1.063 | 1.0662 | 1.094 | 1.122 | 1.046 | 1.01 | 1.057 | 9
| 1 | 6 | 1.0808 | 1.1 | 1.08 |  | 1.1528 | 1.081 | 1.0488 | 1.149 | 1.08 | 1.05 | 1.015 | 1.152
| 2 | 7 | 1.143 | 1.01 | 1.16 |  | 1.05 | 1.07 | 1.088 | 0.991 | 1.084 | 1.055 | 1.0394 | 1.09
| 2 | 8 | 1.0642 | 1.14 | 1.01 |  | 0.99 | 1.04 | 1.054 | 0.977 | 1.095 | 1.132 | 1.1234 | 1.055
| 2 | 9 | 1.0695 | 1.1 | 1.09 |  | 1.11 | 1.11 | 1.075 | 1.119 | 1.187 | 1.095 | 1.0583 | 1.112
| 2 | 10 | 1.1935 | 1.06 | 1.18 |  | 1.1 | 1.04 | 1.072 | 1.066 | 1.12 | 1.129 | 1.1033 | 1.014
| 2 | 11 | 1.0687 | 1.11 | 1.06 |  | 1.05 | 1.09 | 1.045 | 0.983 | 1.066 | 1.011 | 0.97845 | 1.125
| 2 | 12 | 1.082 | 1.11 | 1.16 |  | 1.12 | 1.04 | 1.124 | 0.982 | 1.109 | 1.079 | 1.1164 | 1.014
| 3 | 13 | 2.00: 1.0627; 5.01: 1.1; 5.02: 1.12
| 3 | 14 | 2.00: 1.0707; 5.01: 1.08; 5.02: 1.05
| 3 | 15 | 2.00: 1.1245; 5.01: 1.05; 5.02: 0.97
| 3 | 16 | 2.00: 1.0924; 5.01: 1.07; 5.02: 1.1
| 3 | 17 | 2.00: 1.0923; 5.01: 1.09; 5.02: 1.02
| 3 | 18 | 2.00: 1.1099; 5.01: 1.06; 5.02: 1.13
| 4 | 19 | 2.00: 1.2162; 5.01: 1.11; 5.02: 1.08; 5.03: 1.04
| 4 | 20 | 2.00: 1.0665; 5.01: 1.1; 5.02: 1.06; 5.03: 1
| 4 | 21 | 5.01: 1.11; 5.02: 1.08; 5.03: 1.01
| 4 | 22 | 5.01: 1.04; 5.02: 1.07; 5.03: 1.06
| 4 | 23 | 5.01: 1.02; 5.02: 1.09
| 4 | 24 | 5.01: 1; 5.02: 1.08
| 5 | 25 | 5.02: 1.07; 5.03: 1.08
| 5 | 26 | 5.02: 1.1; 5.03: 1.02
| 5 | 27 | 5.02: 1.08; 5.03: 1.07
| 5 | 28 | 5.02: 1.06; 5.03: 1
| 5 | 29 | 5.02: 1.05
| 5 | 30 | 5.02: 1.04
| 6 | 31 | 5.02: 1.12; 5.03: 1.05
| 6 | 32 | 5.02: 1.07; 5.03: 1.06
| 6 | 33 | 5.02: 1.09; 5.03: 1.04
| 6 | 34 | 5.02: 1.08; 5.03: 0.96
| 6 | 35 | 5.02: 1.11
| 6 | 36 | 5.02: 1.12
| Mean |  | 1.1042 | 1.0771 | 1.0831 | 1.0325 | 1.0659 | 1.0637 | 1.0624 | 1.0415 | 1.0988 | 1.0859 | 1.0561 | 1.0795
| Median |  | 1.0832 | 1.085 | 1.08 | 1.04 | 1.0653 | 1.0645 | 1.0635 | 1.027 | 1.1005 | 1.087 | 1.0489 | 1.0755
| Std Dev. |  | 0.04533 | 0.037005 | 0.042482 | 0.035194 | 0.052433 | 0.022721 | 0.029563 | 0.058782 | 0.038856 | 0.041252 | 0.049756 | 0.043542
| Rel.Std.Dev. |  | 4.105% | 3.436% | 3.922% | 3.409% | 4.919% | 2.136% | 2.783% | 5.644% | 3.536% | 3.799% | 4.711% | 4.034%
| PDM3 |  | 3.21% | 0.6764% | 1.235% | -3.491% | -0.3707% | -0.5777% | -0.6929% | -2.65% | 2.709% | 1.502% | -1.284% | 0.9023%
| Z-Score (Absolute) |  | NA | 0.55 | 0.85 | 1.65 | 0 | 0.11 | 0.17 | 1.21 | 1.63 | 0.99 | 0.48 | 0.67
[a]: Lab 5.02: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (350g), Rnd 03 - Raw*PA (350g), Rnd 04 - Raw*PA (400g), Rnd 05 - Raw*PA (400g), Rnd 06 - Raw*PA (400g)
[b]: Lab 5.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[c]: Lab 5.05: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (300g)
[d]: Lab 5.07: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[e]: Lab 6.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 233b (Certified Value 1.01 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.33 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.085g | 30g | 25g | 25g | 30g | 50g | 30g | 15g | 50g | 30g | 10g | 30g | 25g | 25g | 25g | 25g | 15g | 25g | 15g | 50g | 25g | 25g | 10g | 20g | 3
| 1 | 1 | 1.077 | 0.98 | 1.15 | 1.04 | 1.076 | 0.98 | 0.955 | 1.05 | 0.99 | 0.99 | 0.975 | 1.02 | 1.02 |  |  |  | 0.9999 | 1.038 | 1.0063 | 1.0044 | 1.055 | 0.981 | 1.062 | 1.014 | 1
| 1 | 2 | 1.1651 | 1.03 | 1.13 | 1.04 | 1.109 | 1 | 0.946 | 1.12 | 0.92 | 0.99 | 0.992 | 1.02 | 1.02 | 1.01 | 1 | 1.06 | 0.9797 | 1.012 | 0.9981 | 0.9987 | 1.05 | > 1 | 1.071 | 1.0201 | 0
| 1 | 3 | 1.0784 | 0.99 | 1.14 | 1.12 | 1.111 | 1.04 | 0.984 | 1.17 | 0.96 | 1.02 | 1.046 | 1 |  |  | 1.01 |  | 1.0302 | 1.015 | 0.9985 | 1.0026 | 1.08 | 0.999 | 1.073 | 1.0004 | 16
| 1 | 4 | 1.0844 | 1.1 | 1.13 | 1.12 | 1.122 | 1.01 | 1.01 | 1.14 | 0.92 | 1.03 | 1.015 | 1.052 |  |  |  | 1 | 0.9595 | 1.044 | 1.0488 | 1.0591 | 1.055 | 0.951 | 1.059 | 1.037 | 1.0124
| 1 | 5 | 1.1418 | 1.03 | 1.15 | 1.04 | 1.111 | 1 | 0.926 | 1.08 | 0.96 | 0.99 | 1.047 | 1.036 | 1.01 |  |  |  | 0.9999 | 1.019 | 0.9845 | 1.0581 | 1.14 | 0.972 | 1.066 | 1.0247 | 11
| 1 | 6 | 1.0808 | 1.01 | 1.16 | 1.04 | 1.089 | 0.97 | 0.96 | 1.13 | 0.96 | 1.02 | 1.015 | 1.02 | 1 | 1 |  | 1.01 | 0.9999 | 0.994 | 1.0033 | 1.0098 | 1 | 0.998 | 1.064 | 1.0258
|  | 7 | 1.143
|  | 8 | 1.0642
|  | 9 | 1.0695
|  | 10 | 1.1935
|  | 11 | 1.0687
|  | 12 | 1.082
|  | 13 | 1.0627
|  | 14 | 1.0707
|  | 15 | 1.1245
|  | 16 | 1.0924
|  | 17 | 1.0923
|  | 18 | 1.1099
|  | 19 | 1.2162
|  | 20 | 1.0665
| Mean |  | 1.1042 | 1.0233 | 1.1433 | 1.0667 | 1.103 | 1 | 0.9635 | 1.115 | 0.95167 | 1.0067 | 1.015 | 1.0247 | 1.0125 | 1.005 | 1.005 | 1.0233 | 0.99485 | 1.0203 | 1.0066 | 1.0221 | 1.0633 | 0.9802 | 1.0658 | 1.0203
| Median |  | 1.0832 | 1.02 | 1.145 | 1.04 | 1.11 | 1 | 0.9575 | 1.125 | 0.96 | 1.005 | 1.015 | 1.02 | 1.015 | 1.005 | 1.005 | 1.01 | 0.9999 | 1.017 | 1.0009 | 1.0071 | 1.055 | 0.981 | 1.065 | 1.0224
| Std Dev. |  | 0.04533 | 0.04274 | 0.012111 | 0.041312 | 0.017029 | 0.024495 | 0.029596 | 0.043243 | 0.027142 | 0.018619 | 0.028684 | 0.017603 | 0.0095743 | 0.0070711 | 0.0070711 | 0.032146 | 0.023687 | 0.01825 | 0.021991 | 0.028486 | 0.045789 | 0.019942 | 0.0053448 | 0.012363
| Rel.Std.Dev. |  | 4.105% | 4.177% | 1.059% | 3.873% | 1.544% | 2.449% | 3.072% | 3.878% | 2.852% | 1.85% | 2.826% | 1.718% | 0.9456% | 0.7036% | 0.7036% | 3.141% | 2.381% | 1.789% | 2.185% | 2.787% | 4.306% | 2.035% | 0.5015% | 1.212%
| PDM3 |  | 9.071% | 1.084% | 12.94% | 5.365% | 8.954% | -1.221% | -4.826% | 10.14% | -5.995% | -0.5621% | 0.2611% | 1.216% | 0.01413% | -0.7267% | -0.7267% | 1.084% | -1.729% | 0.7879% | -0.5703% | 0.9641% | 5.035% | -3.176% | 5.282% | 0.7879%
| Z-Score (Absolute) |  | NA | 0.13 | 5.41 | 2.04 | 3.64 | 0.89 | 2.5 | 4.16 | 3.02 | 0.6 | 0.23 | 0.19 | 0.34 | 0.67 | 0.67 | 0.13 | 1.12 | 0 | 0.6 | 0.08 | 1.89 | 1.76 | 2 | 0
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 233b (Certified Value 1.01 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.30 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 05g | 50g | 60g | 30g | 30g | 30g | 200g | 30g | 200g | 200g | 20g | 30g | 05g | 3
| 1 | 1 | 1.077 | 0.917 | 1.06 | 0.975 | 0.968 | 0.94458 | 0.928 | 0.986 | 1.0416 | 1 | 1.0037 | 0.99 | 1.11 | 1.05 | 1.07 | 1.013 | 1.07 | 0.967 | 1.0441 | 1.01 | 1.02 | 1.04 | 1
| 1 | 2 | 1.1651 | 0.911 | 1.05 | 0.875 | 0.936 | 0.9415 | 0.988 | 0.957 | 1.0416 | 1 | 1 | 0.95 | 1.12 | 0.98 | 1.05 | 1.063 | 1.05 | 0.967 | 1.0365 | 0.97 | 1.009 | 1.04 | 0
| 1 | 3 | 1.0784 | 0.962 | 1.03 | 0.925 | 1.056 | 0.90695 | 0.984 | 0.964 | 1.08 | 1 | 1.038 | 0.98 | 1.09 | 0.98 | 0.94 | 1.043 | 1.08 | 0.959 | 1.0278 | 1 | 1.062 | 1.05 | 16
| 1 | 4 | 1.0844 | 0.941 | 1.03 | 0.9 | 0.984 | 0.92171 | 0.933 | 1 | 1.04 | 1 | 1.026 | 0.97 | 1.14 | 0.92 | 0.95 | 1.033 | 1.08 | 0.962 | 1.0503 | 0.97 | 1.061 | 1.05 | 1.0053
| 1 | 5 | 1.1418 | 0.943 | 1.01 | 0.975 | 1.024 | 0.90709 | 1.01 | 0.977 | 1.04 | 1 | 1.0299 | 1 | 1.13 | 0.92 | 1.06 | 1.07 | 1.1 | 0.964 | 1.0321 | 1.01 | 1.066 | 1.02 | 13
| 1 | 6 | 1.0808 | 0.884 | 1.03 | 1 | 0.992 | 0.91531 | 0.975 | 0.98 | 1.04 | 1.2 | 1.028 | 0.98 | 1.12 | 0.94 | 1.05 | 1.033 | 1.08 | 0.962 | 1.0152 | 1 | 1.068 | 1.03
|  | 7 | 2.00: 1.143
|  | 8 | 2.00: 1.0642
|  | 9 | 2.00: 1.0695
|  | 10 | 2.00: 1.1935
|  | 11 | 2.00: 1.0687
|  | 12 | 2.00: 1.082
|  | 13 | 2.00: 1.0627
|  | 14 | 2.00: 1.0707
|  | 15 | 2.00: 1.1245
|  | 16 | 2.00: 1.0924
|  | 17 | 2.00: 1.0923
|  | 18 | 2.00: 1.1099
|  | 19 | 2.00: 1.2162
|  | 20 | 2.00: 1.0665
| Mean |  | 1.1042 | 0.92633 | 1.035 | 0.94167 | 0.99333 | 0.92286 | 0.96967 | 0.97733 | 1.0472 | 1.0333 | 1.0209 | 0.97833 | 1.1183 | 0.965 | 1.02 | 1.0425 | 1.0767 | 0.9635 | 1.0343 | 0.99333 | 1.0477 | 1.0383
| Median |  | 1.0832 | 0.929 | 1.03 | 0.95 | 0.988 | 0.91851 | 0.9795 | 0.9785 | 1.0408 | 1 | 1.027 | 0.98 | 1.12 | 0.96 | 1.05 | 1.038 | 1.08 | 0.963 | 1.0343 | 1 | 1.0615 | 1.04
| Std Dev. |  | 0.04533 | 0.027854 | 0.017607 | 0.04916 | 0.042155 | 0.016609 | 0.032488 | 0.01541 | 0.016082 | 0.08165 | 0.015387 | 0.017224 | 0.017224 | 0.049699 | 0.058652 | 0.021107 | 0.01633 | 0.0031464 | 0.012401 | 0.018619 | 0.026051 | 0.01169
| Rel.Std.Dev. |  | 4.105% | 3.007% | 1.701% | 5.22% | 4.244% | 1.8% | 3.35% | 1.577% | 1.536% | 7.902% | 1.507% | 1.761% | 1.54% | 5.15% | 5.75% | 2.025% | 1.517% | 0.3266% | 1.199% | 1.874% | 2.487% | 1.126%
| PDM3 |  | 9.831% | -7.86% | 2.949% | -6.334% | -1.195% | -8.205% | -3.549% | -2.787% | 4.164% | 2.783% | 1.55% | -2.687% | 11.24% | -4.014% | 1.457% | 3.695% | 7.094% | -4.163% | 2.884% | -1.195% | 4.209% | 3.281%
| Z-Score (Absolute) |  | NA | 2.28 | 0.37 | 1.91 | 0.65 | 2.37 | 1.23 | 1.04 | 0.66 | 0.32 | 0.02 | 1.02 | 2.4 | 1.34 | 0 | 0.55 | 1.38 | 1.38 | 0.35 | 0.65 | 0.67 | 0.45
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 233b (Certified Value 0.311 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.45 | 0.31 | 0.29935 | < 0.3 | 0.46 | 0.4 | 0.28 | < 0.5 | 0.3 | 0.32 | 0.31 | 0.32 | 0.3 | < 0.2 | 0.3 | 0.3 | 0.31 | 0.25 | < 1 | 0.3 | < 0.3 | 0.31 | 0.31 | 0.35 | 0.27 |  |  |  |  | 1
| 1 | 2 | 0.43 | 0.29 | 0.30946 | < 0.3 | 0.43 | 0.3 | 0.31 | < 0.5 | 0.3 | 0.34 | 0.3 | 0.31 | 0.3 | < 0.2 | 0.3 | 0.3 | 0.33 | 0.22 | < 1 | 0.31 | 0.3102 | 0.29 | 0.32 | 0.35 | 0.27 |  |  |  |  | 18
| 1 | 3 | 0.45 | 0.33 | 0.30228 | 0.3 | 0.42 | 0.3 | 0.31 | < 0.5 | 0.3 | 0.32 | 0.31 | 0.33 | 0.31 | 0.2 | 0.3 | 0.3 | 0.34 | 0.32 | < 1 | 0.32 | < 0.3 | 0.3 | 0.3 | 0.34 | 0.25 |  |  |  |  | 16
| 1 | 4 | 0.42 | 0.32 | 0.31823 | < 0.3 | 0.43 | 0.3 | 0.3 | < 0.5 | 0.3 | 0.33 | 0.32 | 0.31 | 0.31 | 0.2 | 0.3 | 0.3 | 0.34 | 0.32 | < 1 | 0.32 | < 0.3 | 0.3 | 0.29 | 0.36 | 0.25 |  |  |  |  | 0.31065
| 1 | 5 | 0.42 | 0.31 | 0.31288 | < 0.3 | 0.41 | 0.3 | 0.3 | < 0.5 | 0.4 | 0.32 | 0.31 | 0.31 | 0.29 | 0.3 | 0.3 | 0.3 | 0.34 | 0.43 | < 1 | 0.3 | < 0.3 | 0.33 | 0.3 | 0.36 | 0.27 |  |  |  |  | 15
| 1 | 6 | 0.45 | 0.32 | 0.32703 | < 0.3 | 0.41 | 0.3 | 0.29 | < 0.5 | 0.3 | 0.34 | 0.32 | 0.32 | 0.33 | 0.2 | 0.3 | 0.3 | 0.32 | 0.29 | < 1 | 0.31 | < 0.3 | 0.28 | 0.31 | 0.35 | 0.24
| Mean |  | 0.43667 | 0.31333 | 0.31154 | 0.3 | 0.42667 | 0.31667 | 0.29833 |  | 0.31667 | 0.32833 | 0.31167 | 0.31667 | 0.30667 | 0.225 | 0.3 | 0.3 | 0.33 | 0.305 |  | 0.31 | 0.3102 | 0.30167 | 0.305 | 0.35167 | 0.25833
| Median |  | 0.44 | 0.315 | 0.31117 | 0.3 | 0.425 | 0.3 | 0.3 |  | 0.3 | 0.325 | 0.31 | 0.315 | 0.305 | 0.2 | 0.3 | 0.3 | 0.335 | 0.305 |  | 0.31 | 0.3102 | 0.3 | 0.305 | 0.35 | 0.26
| Std Dev. |  | 0.015055 | 0.013663 | 0.01025 |  | 0.018619 | 0.040825 | 0.01169 |  | 0.040825 | 0.0098319 | 0.0075277 | 0.008165 | 0.013663 | 0.05 | 0 | 0 | 0.012649 | 0.07287 |  | 0.0089443 |  | 0.017224 | 0.010488 | 0.0075277 | 0.013292
| Rel.Std.Dev. |  | 3.448% | 4.36% | 3.29% |  | 4.364% | 12.89% | 3.919% |  | 12.89% | 2.994% | 2.415% | 2.578% | 4.455% | 22.22% | 0% | 0% | 3.833% | 23.89% |  | 2.885% |  | 5.71% | 3.439% | 2.141% | 5.145%
| PDM3 |  | 40.57% | 0.8649% | 0.2872% | -3.427% | 37.35% | 1.938% | -3.964% |  | 1.938% | 5.694% | 0.3284% | 1.938% | -1.281% | -27.57% | -3.427% | -3.427% | 6.23% | -1.818% |  | -0.2081% | -0.1438% | -2.891% | -1.818% | 13.2% | -16.84%
| Z-Score (Absolute) |  | 5.66 | 0.12 | 0.03 | NA | 5.21 | NA | 0.56 | 2.73 | NA | 0.79 | 0.04 | 0.27 | 0.18 | 5.73 | NA | NA | 0.86 | 0.26 | 8.51 | 0.03 | 6.03 | 0.41 | 0.26 | 1.84 | 2.36
Results from laboratories 1.07, 1.10, 1.13, 1.22 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 233b (Certified Value 6.8 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 7.35 | 6.6478 | 6.9373 | 6.7428 | 6.66 | 7.568 | 7.04 | 6.57 | 6.62 | 6.51 | 6.96 | 6.72 | 6.58 | 6.61 | 6.67 | 6.848 | 6.74 | 6.74 | 3.85 | 6.3964 | 6.72 | 6.86 | 6.157 | 6.8 | 7.4 | 7.03 | 7.1 |  |  | 1
| 1 | 2 | 7.46 | 6.5749 | 6.9056 | 6.8422 | 6.67 | 7.568 | 6.85 | 6.6232 | 6.63 | 6.76 | 7.19 | 6.52 | 6.78 | 6.58 | 6.89 | 6.876 | 6.82 | 6.77 | 3.99 | 6.4147 | 6.701 | 6.81 | 6.483 | 6.9 | 7.23 | 7.12 | 7.2 |  |  | 0
| 1 | 3 | 7.33 | 6.6631 | 6.9875 | 6.7644 | 6.59 | 7.515 | 6.98 | 6.5522 | 6.61 | 6.36 | 7.18 | 6.46 | 6.59 | 6.64 | 6.8 | 6.783 | 6.82 | 6.61 | 3.83 | 6.6584 | 6.706 | 6.82 | 6.369 | 6.72 | 7.3 | 7.02 | 7.2 |  |  | 16
| 1 | 4 | 7.15 | 6.5009 | 6.9204 | 6.8626 | 6.65 | 7.515 | 6.99 | 6.6241 | 6.6 | 6.79 | 7.16 | 6.5 | 6.54 | 6.65 | 6.81 | 6.957 | 6.74 | 6.55 | 4.01 | 6.5187 | 6.743 | 6.72 | 6.362 | 6.93 | 7.23 | 7.04 | 7.3 |  |  | 6.7985
| 1 | 5 | 7.14 | 6.5051 | 6.9647 | 6.9728 | 6.63 | 7.462 | 6.98 | 6.411 | 6.64 | 6.83 | 7.18 | 6.66 | 6.46 | 6.55 | 6.95 | 6.833 | 6.79 | 7.09 | 3.49 | 6.4505 | 6.691 | 6.72 | 6.208 | 6.87 | 7.19 | 7.05 | 7.2 |  |  | 16
| 1 | 6 | 7.34 | 6.5389 | 6.9731 | 6.4704 | 6.56 | 7.568 | 7.17 | 6.5762 | 6.63 | 6.73 | 7.15 | 6.61 | 6.98 | 6.7 | 6.78 | 6.922 | 6.79 | 6.68 | 3.4 | 6.649 | 6.741 | 6.82 | 6.45 | 6.88 | 7.27 | 7.03 | 7.4
| Mean |  | 7.295 | 6.5718 | 6.9481 | 6.7759 | 6.6267 | 7.5327 | 7.0017 | 6.5595 | 6.6217 | 6.6633 | 7.1367 | 6.5783 | 6.655 | 6.6217 | 6.8167 | 6.8698 | 6.7833 | 6.74 | 3.7617 | 6.5146 | 6.717 | 6.7917 | 6.3382 | 6.85 | 7.27 | 7.0483 | 7.2333
| Median |  | 7.335 | 6.5569 | 6.951 | 6.8033 | 6.64 | 7.5415 | 6.985 | 6.5731 | 6.625 | 6.745 | 7.17 | 6.565 | 6.585 | 6.625 | 6.805 | 6.862 | 6.79 | 6.71 | 3.84 | 6.4846 | 6.713 | 6.815 | 6.3655 | 6.875 | 7.25 | 7.035 | 7.2
| Std Dev. |  | 0.12534 | 0.070256 | 0.032065 | 0.17045 | 0.043205 | 0.043274 | 0.10381 | 0.078391 | 0.01472 | 0.18608 | 0.087788 | 0.10128 | 0.19097 | 0.053448 | 0.096264 | 0.062819 | 0.036148 | 0.18974 | 0.25725 | 0.1156 | 0.021513 | 0.058109 | 0.13021 | 0.076942 | 0.074027 | 0.03656 | 0.10328
| Rel.Std.Dev. |  | 1.718% | 1.069% | 0.4615% | 2.516% | 0.652% | 0.5745% | 1.483% | 1.195% | 0.2223% | 2.793% | 1.23% | 1.54% | 2.87% | 0.8072% | 1.412% | 0.9144% | 0.5329% | 2.815% | 6.839% | 1.774% | 0.3203% | 0.8556% | 2.054% | 1.123% | 1.018% | 0.5187% | 1.428%
| PDM3 |  | 7.303% | -3.335% | 2.2% | -0.3331% | -2.528% | 10.8% | 2.988% | -3.517% | -2.601% | -1.989% | 4.974% | -3.239% | -2.111% | -2.601% | 0.2668% | 1.049% | -0.2235% | -0.8609% | -44.67% | -4.176% | -1.199% | -0.1009% | -6.771% | 0.7571% | 6.935% | 3.674% | 6.396%
| Z-Score (Absolute) |  | 2.03 | 0.8 | 0.67 | 0 | 0.58 | 2.96 | 0.88 | 0.85 | 0.6 | 0.44 | 1.41 | 0.77 | 0.47 | 0.6 | 0.16 | 0.37 | 0.03 | 0.14 | 11.8 | 1.02 | 0.23 | 0.06 | 1.71 | 0.29 | 1.93 | 1.07 | 1.79
| Analytical results for As in OREAS 233b (Certified Value 60 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 69 | 61.9 | 63.5398 | 55.6 | 65 | 59 | 56 | 47.8 | 60 | 55.8 | 59.4 | 58.1 | 58 | 61 | 57 | 59 | 34.42 | 54 | 60.6 | 58.7112 | 61.6 | 64 | 59 | 54.3 | [5 empty] | 1
| 1 | 2 | 69 | 61.2 | 64.1953 | 47.8 | 61 | 59.9 | 55 | 45.4 | 63.3 | 54.2 | 61.7 | 60.7 | 61 | 57 | 57 | 60 | 31.66 | 56 | 61 | 58.5094 | 61.2 | 62 | 58.4 | 52.2 | [5 empty] | 19
| 1 | 3 | 69 | 61 | 63.1516 | 55.8 | 63 | 59 | 55 | 47.3 | 62.6 | 58 | 59.1 | 60.8 | 61 | 57 | 58 | 58 | 41.29 | 57 | 59.9 | 58.7006 | 62.6 | 61 | 60.4 | 49.9 | [5 empty] | 16
| 1 | 4 | 67 | 60.2 | 63.4407 | 44.3 | 62 | 60.1 | 56 | 44 | 61.2 | 55.5 | 58.8 | 60.3 | 60 | 60 | 57 | 59 | 37.772 | 58 | 59.9 | 58.5074 | 63.2 | 61 | 61.2 | 49.5 | [5 empty] | 59.856
| 1 | 5 | 67 | 61.4 | 65.0906 | 45.8 | 61 | 58.2 | 56 | 45.5 | 61.4 | 57.6 | 61.1 | 61.7 | 61 | 60 | 57 | 64 | 50.21 | 56 | 63.2 | 57.6319 | 64.1 | 58 | 61.9 | 50.9 | [5 empty] | 17
| 1 | 6 | 70 | 61.2 | 61.0487 | 38.4 | 64 | 59.3 | 56 | 46.9 | 61.6 | 61.3 | 60.9 | 61.2 | 58 | 64 | 57 | 59 | 38.497 | 57 | 61.7 | 60.368 | 62.1 | 60 | 60.1 | 48
| Mean |  | 68.5 | 61.15 | 63.4111 | 47.95 | 62.6667 | 59.25 | 55.6667 | 46.15 | 61.6833 | 57.0667 | 60.1667 | 60.4667 | 59.8333 | 59.8333 | 57.1667 | 59.8333 | 38.9748 | 56.3333 | 61.05 | 58.7381 | 62.4667 | 61 | 60.1667 | 50.8
| Median |  | 69 | 61.2 | 63.4903 | 46.8 | 62.5 | 59.15 | 56 | 46.2 | 61.5 | 56.7 | 60.15 | 60.75 | 60.5 | 60 | 57 | 59 | 38.1345 | 56.5 | 60.8 | 58.605 | 62.35 | 61 | 60.25 | 50.4
| Std Dev. |  | 1.2247 | 0.55767 | 1.3492 | 6.7722 | 1.633 | 0.6892 | 0.5164 | 1.4293 | 1.1496 | 2.5041 | 1.2127 | 1.2533 | 1.472 | 2.6394 | 0.40825 | 2.137 | 6.4423 | 1.3663 | 1.2566 | 0.89307 | 1.069 | 2 | 1.3125 | 2.2163
| Rel.Std.Dev. |  | 1.788% | 0.912% | 2.128% | 14.12% | 2.606% | 1.163% | 0.9277% | 3.097% | 1.864% | 4.388% | 2.016% | 2.073% | 2.46% | 4.411% | 0.7141% | 3.572% | 16.53% | 2.425% | 2.058% | 1.52% | 1.711% | 3.279% | 2.181% | 4.363%
| PDM3 |  | 14.44% | 2.162% | 5.94% | -19.89% | 4.696% | -1.012% | -6.999% | -22.9% | 3.053% | -4.66% | 0.5191% | 1.02% | -0.0378% | -0.0378% | -4.493% | -0.0378% | -34.89% | -5.885% | 1.995% | -1.868% | 4.362% | 1.911% | 0.5191% | -15.13%
| Z-Score (Absolute) |  | 2.61 | 0.4 | 1.08 | 3.57 | 0.85 | 0.18 | 1.25 | 4.12 | 0.56 | 0.83 | 0.1 | 0.19 | 0 | 0 | 0.8 | 0 | 6.27 | 1.05 | 0.37 | 0.33 | 0.79 | 0.35 | 0.1 | 2.72
| Analytical results for B in OREAS 233b (Indicative Value 36.8 ppm) | [30 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.02: 1.25; 1.04: 1.27 |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES/MS; 1.04: 4A*OES/MS |  |  | 1
1
| 1 | 1 | 1.01: 53.24; 1.02: 13547.08; 1.04: 10 |  |  | 1
| 1 | 2 | 1.01: 55.83; 1.02: 13698.79; 1.04: 10 |  |  | 3
| 1 | 3 | 1.01: 60.76; 1.02: 20341.49; 1.04: 10 |  |  | 16
| 1 | 4 | 1.01: 76.64; 1.02: 16306.08; 1.04: 10 |  |  | 36.815
| 1 | 5 | 1.01: 58.68; 1.02: 32181.58; 1.04: 10 |  |  | 9
| 1 | 6 | 1.01: 76.63; 1.02: 19233.86; 1.04: 10
| Mean |  | 1.01: 63.63; 1.02: 19218.1467; 1.04: 10
| Median |  | 1.01: 59.72; 1.02: 17769.97; 1.04: 10
| Std Dev. |  | 1.01: 10.3903; 1.02: 6934.2221; 1.04: 0
| Rel.Std.Dev. |  | 1.01: 16.33%; 1.02: 36.08%; 1.04: 0%
| PDM3 |  | 1.01: 72.84%; 1.02: 52102%; 1.04: -72.84%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 240.84; 1.04: 0.67
| Analytical results for Ba in OREAS 233b (Certified Value 270 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 277 | 281.9 | 282.4157 | 264.4357 | 272 | 307 | 278 | 273.3 | 265 | 272 | 290 | 280 | 260 | 270 | 213 | 268 | 254 | 253 | 517.08 | 248.634 | 270 | 268.9923 | 280 | 297.7405 | 274 | 316 | 263 | 256 |  | 1
| 1 | 2 | 278 | 275.8 | 285.0628 | 270.1107 | 259 | 313 | 272 | 273.8 | 265 | 274 | 300 | 280 | 280 | 270 | 218 | 272 | 250 | 255 | 838.39 | 250.623 | 270 | 271.7222 | 280 | 299.7775 | 278 | 306 | 267 | 239 |  | 20
| 1 | 3 | 276 | 274.9 | 284.9186 | 266.2183 | 261 | 309 | 274 | 270.9 | 261 | 267 | 300 | 270 | 260 | 270 | 215 | 265 | 247 | 247 | 1061.02 | 261.846 | 270 | 273.8132 | 280 | 303.0688 | 272 | 302 | 264 | 231 |  | 16
| 1 | 4 | 272 | 273.3 | 283.8989 | 269.4976 | 252 | 311 | 276 | 273.7 | 264 | 277 | 300 | 270 | 260 | 270 | 217 | 272 | 256 | 245 | 975.6 | 255.744 | 270 | 266.0269 | 270 | 300.5376 | 280 | 306 | 266 | 228 |  | 270.0018
| 1 | 5 | 268 | 270.2 | 288.6369 | 277.7477 | 255 | 305 | 271 | 266.1 | 260 | 276 | 290 | 280 | 260 | 270 | 226 | 267 | 255 | 265 | 1116.74 | 252.792 | 270 | 263.0016 | 280 | 294.4659 | 277 | 303 | 265 | 237 |  | 18
| 1 | 6 | 275 | 279.9 | 281.911 | 256.4537 | 256 | 310 | 279 | 275.1 | 261 | 268 | 290 | 270 | 280 | 280 | 216 | 271 | 253 | 250 | 1038.52 | 260.172 | 275 | 268.6564 | 280 | 293.9607 | 278 | 308 | 263 | 224
| Mean |  | 274.3333 | 276 | 284.474 | 267.4106 | 259.1667 | 309.1667 | 275 | 272.15 | 262.6667 | 272.3333 | 295 | 275 | 266.6667 | 271.6667 | 217.5 | 269.1667 | 252.5 | 252.5 | 924.5583 | 254.9685 | 270.8333 | 268.7021 | 278.3333 | 298.2585 | 276.5 | 306.8333 | 264.6667 | 235.8333
| Median |  | 275.5 | 275.35 | 284.4087 | 267.858 | 257.5 | 309.5 | 275 | 273.5 | 262.5 | 273 | 295 | 275 | 260 | 270 | 216.5 | 269.5 | 253.5 | 251.5 | 1007.06 | 254.268 | 270 | 268.8243 | 280 | 298.759 | 277.5 | 306 | 264.5 | 234
| Std Dev. |  | 3.7238 | 4.2942 | 2.4088 | 7.0545 | 7.0261 | 2.8577 | 3.2249 | 3.2654 | 2.2509 | 4.1312 | 5.4772 | 5.4772 | 10.328 | 4.0825 | 4.5056 | 2.9269 | 3.3912 | 7.1484 | 221.1839 | 5.2673 | 2.0412 | 3.8725 | 4.0825 | 3.5711 | 2.9496 | 4.9967 | 1.633 | 11.3387
| Rel.Std.Dev. |  | 1.357% | 1.556% | 0.8467% | 2.638% | 2.711% | 0.9243% | 1.173% | 1.2% | 0.857% | 1.517% | 1.857% | 1.992% | 3.873% | 1.503% | 2.072% | 1.087% | 1.343% | 2.831% | 23.92% | 2.066% | 0.7537% | 1.441% | 1.467% | 1.197% | 1.067% | 1.628% | 0.617% | 4.808%
| PDM3 |  | 1.604% | 2.222% | 5.36% | -0.9597% | -4.013% | 14.51% | 1.851% | 0.7956% | -2.717% | 0.8635% | 9.259% | 1.851% | -1.235% | 0.6166% | -19.44% | -0.3093% | -6.482% | -6.482% | 242.4% | -5.568% | 0.308% | -0.4813% | 3.086% | 10.47% | 2.407% | 13.64% | -1.976% | -12.65%
| Z-Score (Absolute) |  | 0.24 | 0.4 | 1.24 | 0.44 | 1.26 | 3.68 | 0.31 | 0.02 | 0.91 | 0.04 | 2.28 | 0.31 | 0.52 | 0.02 | 5.37 | 0.27 | 1.92 | 1.92 | 64.4 | 1.67 | 0.11 | 0.32 | 0.63 | 2.6 | 0.45 | 3.45 | 0.71 | 3.56
| Analytical results for Be in OREAS 233b (Certified Value 0.39 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.4 | 0.46 | 0.22 | 0.36435 | < 0.5 | 0.4 | 0.6 | 0.6 | 0.35 | < 5 | < 1 | 0.42 | 0.43 | 0.43 | 0.38 | < 0.5 | < 0.5 | < 1 | 0.37 | 0.5 | 0.34744 | 0.39 | 0.31014 | < 0.05 | 0.5 | 0.44 | 0.32 |  |  | 1
| 1 | 2 | 0.4 | 0.44 | 0.21 | 0.37394 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.43 | < 5 | < 1 | 0.43 | 0.42 | 0.42 | 0.38 | < 0.5 | < 0.5 | < 1 | 0.4 | < 0.5 | 0.33314 | 0.38 | 0.31971 | < 0.05 | 0.5 | 0.45 | 0.32 |  |  | 21
| 1 | 3 | 0.4 | 0.43 | 0.22 | 0.38832 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.33 | < 5 | < 1 | 0.45 | 0.43 | 0.39 | 0.38 | < 0.5 | < 0.5 | < 1 | 0.4 | < 0.5 | 0.33114 | 0.4 | 0.31647 | < 0.05 | 0.5 | 0.46 | 0.33 |  |  | 16
| 1 | 4 | 0.4 | 0.42 | 0.38 | 0.4057 | < 0.5 | 0.4 | 0.5 | < 0.5 | 0.35 | < 5 | < 1 | 0.44 | 0.43 | 0.43 | 0.38 | < 0.5 | < 0.5 | < 1 | 0.41 | < 0.5 | 0.37048 | 0.38 | 0.3166 | < 0.05 | 0.5 | 0.47 | 0.33 |  |  | 0.39138
| 1 | 5 | 0.4 | 0.48 | 0.34 | 0.40211 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.42 | < 5 | < 1 | 0.44 | 0.43 | 0.42 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.36 | 0.5 | 0.35266 | 0.39 | 0.3076 | < 0.05 | 0.5 | 0.47 | 0.42 |  |  | 19
| 1 | 6 | 0.4 | 0.45 | 0.34 | 0.38233 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.4 | < 5 | < 1 | 0.43 | 0.43 | 0.46 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.38 | < 0.5 | 0.37593 | 0.38 | 0.31397 | < 0.05 | 0.5 | 0.44 | 0.5
| Mean |  | 0.4 | 0.44667 | 0.285 | 0.38613 |  | 0.4 | 0.51667 | 0.6 | 0.38 |  |  | 0.435 | 0.42833 | 0.425 | 0.38667 |  |  |  | 0.38667 | 0.5 | 0.3518 | 0.38667 | 0.31408 |  | 0.5 | 0.455 | 0.37
| Median |  | 0.4 | 0.445 | 0.28 | 0.38533 |  | 0.4 | 0.5 | 0.6 | 0.375 |  |  | 0.435 | 0.43 | 0.425 | 0.38 |  |  |  | 0.39 | 0.5 | 0.35005 | 0.385 | 0.31522 |  | 0.5 | 0.455 | 0.33
| Std Dev. |  | 6.0809e-17 | 0.021602 | 0.076354 | 0.016003 |  | 6.0809e-17 | 0.040825 | 0 | 0.041952 |  |  | 0.010488 | 0.0040825 | 0.022583 | 0.010328 |  |  |  | 0.019664 | 0 | 0.018579 | 0.008165 | 0.0045013 |  | 0 | 0.013784 | 0.074297
| Rel.Std.Dev. |  | 1.52e-14% | 4.836% | 26.79% | 4.144% |  | 1.52e-14% | 7.902% | 0% | 11.04% |  |  | 2.411% | 0.9531% | 5.314% | 2.671% |  |  |  | 5.085% | 0% | 5.281% | 2.112% | 1.433% |  | 0% | 3.029% | 20.08%
| PDM3 |  | 2.202% | 14.13% | -27.18% | -1.343% |  | 2.202% | 32.01% | 53.3% | -2.908% |  |  | 11.14% | 9.441% | 8.59% | -1.205% |  |  |  | -1.205% | 27.75% | -10.11% | -1.205% | -19.75% |  | 27.75% | 16.25% | -5.463%
| Z-Score (Absolute) |  | NA | 0.63 | 1.07 | 0.01 | 1.44 | NA | NA | 1.63 | 0.07 | 22.21 | 1.19 | 0.51 | 0.44 | 0.4 | 0 | 1.44 | 1.44 | 1.19 | 0 | NA | 0.37 | 0 | 0.76 | 3.8 | NA | 0.72 | 0.18
Results from laboratories 1.01, 1.07, 1.08, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 233b (Certified Value 0.046 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.06 | 0.05 | 0.066043 | 0.05 | 0.2 | 0.05 | 0.04 | < 0.1 | 0.04 | 0.05 | 0.04 | 0.03 | < 5 | < 0.05 | < 0.1 | 0.05 | 0.52 | 0.2 | < 0.11 | 0.04 | 0.0802 | 0.04 | 0.06 | 0.05 | 0.03 |  |  |  |  | 1
| 1 | 2 | 0.06 | 0.04 | 0.064273 | 0.04 | 0.2 | 0.06 | 0.05 | < 0.1 | 0.05 | 0.04 | 0.03 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | 0.1 | < 0.11 | 0.04 | 0.0453 | 0.04 | 0.06 | 0.07 | 0.03 |  |  |  |  | 22
| 1 | 3 | 0.05 | 0.05 | 0.065054 | 0.05 | 0.2 | 0.07 | 0.05 | < 0.1 | 0.04 | 0.04 | 0.04 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.04 | 0.77 | < 0.1 | < 0.11 | 0.04 | 0.0384 | 0.04 | 0.04 | 0.05 | 0.02 |  |  |  |  | 16
| 1 | 4 | 0.06 | 0.05 | 0.067183 | 0.04 | 0.2 | 0.06 | 0.06 | < 0.1 | 0.04 | 0.05 | 0.04 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.04 | 0.0203 | 0.04 | 0.04 | 0.05 | 0.02 |  |  |  |  | 0.045568
| 1 | 5 | 0.06 | 0.04 | 0.067004 | 0.05 | 0.2 | 0.06 | 0.05 | < 0.1 | 0.04 | 0.04 | 0.04 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.04 | 0.0183 | 0.05 | 0.04 | 0.05 | 0.02 |  |  |  |  | 20
| 1 | 6 | 0.08 | 0.04 | 0.072495 | 0.04 | 0.2 | 0.07 | 0.05 | < 0.1 | 0.04 | 0.05 | 0.04 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.05 | 0.18 | < 0.1 | < 0.11 | 0.04 | 0.022 | 0.05 | 0.04 | 0.05 | 0.03
| Mean |  | 0.061667 | 0.045 | 0.067009 | 0.045 | 0.2 | 0.061667 | 0.05 |  | 0.041667 | 0.045 | 0.038333 | 0.038333 |  |  |  | 0.048333 | 0.49 | 0.15 |  | 0.04 | 0.037417 | 0.043333 | 0.046667 | 0.053333 | 0.025
| Median |  | 0.06 | 0.045 | 0.066524 | 0.045 | 0.2 | 0.06 | 0.05 |  | 0.04 | 0.045 | 0.04 | 0.04 |  |  |  | 0.05 | 0.52 | 0.15 |  | 0.04 | 0.0302 | 0.04 | 0.04 | 0.05 | 0.025
| Std Dev. |  | 0.0098319 | 0.0054772 | 0.0029109 | 0.0054772 | 3.0405e-17 | 0.0075277 | 0.0063246 |  | 0.0040825 | 0.0054772 | 0.0040825 | 0.0040825 |  |  |  | 0.0040825 | 0.29614 | 0.070711 |  | 0 | 0.02362 | 0.005164 | 0.010328 | 0.008165 | 0.0054772
| Rel.Std.Dev. |  | 15.94% | 12.17% | 4.344% | 12.17% | 1.52e-14% | 12.21% | 12.65% |  | 9.798% | 12.17% | 10.65% | 10.65% |  |  |  | 8.447% | 60.44% | 47.14% |  | 0% | 63.13% | 11.92% | 22.13% | 15.31% | 21.91%
| PDM3 |  | 35.33% | -1.247% | 47.05% | -1.247% | 338.9% | 35.33% | 9.725% |  | -8.562% | -1.247% | -15.88% | -15.88% |  |  |  | 6.068% | 975.3% | 229.2% |  | -12.22% | -17.89% | -4.905% | 2.41% | 17.04% | -45.14%
| Z-Score (Absolute) |  | 1.52 | 0.17 | 2.06 | 0.17 | NA | 1.52 | 0.34 | 0.34 | 0.51 | 0.17 | 0.84 | 0.84 | 248.15 | 2.19 | 0.34 | 0.17 | 22.59 | NA | 0.84 | 0.67 | 0.94 | 0.34 | 0 | 0.67 | 2.19
Results from laboratories 1.08 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 233b (Certified Value 6.33 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.22 | 6.6399 | 6.4035 | 6.2352 | 6.35 | 7.218 | 6.48 | 6.4689 | 6.13 | 6.23 | 6.5 | 6.36 | 6.28 | 6.34 | 6.07 | 6.571 | 6.06 | 6.408 | 4.28 | 5.8965 | 6.481 | 6.55 | 6.5397 | 6.15 | 5.92 | 6.09 | 6.8 |  |  | 1
| 1 | 2 | 6.2 | 6.5744 | 6.3965 | 6.3311 | 6.14 | 7.147 | 6.32 | 6.6298 | 6.12 | 6.29 | 6.65 | 6.17 | 6.48 | 6.24 | 6.29 | 6.536 | 6.13 | 6.457 | 4.51 | 5.9195 | 6.436 | 6.64 | 6.7397 | 6.25 | 5.82 | 6.2 | 6.8 |  |  | 0
| 1 | 3 | 6.18 | 6.61 | 6.3926 | 6.2978 | 6.11 | 7.147 | 6.41 | 6.4951 | 6.11 | 6.12 | 6.66 | 6.15 | 6.3 | 6.34 | 6.2 | 6.502 | 6.14 | 6.259 | 4.06 | 6.1661 | 6.395 | 6.56 | 6.6827 | 6.06 | 5.88 | 6.13 | 6.9 |  |  | 16
| 1 | 4 | 6.07 | 6.4773 | 6.4438 | 6.391 | 5.95 | 7.147 | 6.42 | 6.4756 | 6.1 | 6.22 | 6.64 | 6.18 | 6.32 | 6.38 | 6.18 | 6.598 | 6.14 | 6.244 | 4.41 | 6.0415 | 6.477 | 6.49 | 6.7357 | 6.24 | 5.81 | 6.16 | 6.9 |  |  | 6.3251
| 1 | 5 | 6.01 | 6.4625 | 6.4745 | 6.4666 | 5.83 | 7.147 | 6.4 | 6.3471 | 6.11 | 6.2 | 6.66 | 6.28 | 6.17 | 6.27 | 6.3 | 6.494 | 6.15 | 6.711 | 3.13 | 5.9473 | 6.497 | 6.53 | 6.5507 | 6.17 | 5.88 | 6.15 | 6.9 |  |  | 21
| 1 | 6 | 6.18 | 6.6166 | 6.4672 | 6.0148 | 5.93 | 7.147 | 6.56 | 6.3052 | 6.12 | 6.15 | 6.6 | 6.37 | 6.7 | 6.27 | 6.07 | 6.59 | 6.08 | 6.316 | 3.88 | 6.1413 | 6.382 | 6.6 | 6.6427 | 6.21 | 5.8 | 6.11 | 6.9
| Mean |  | 6.1433 | 6.5634 | 6.4297 | 6.2894 | 6.0517 | 7.1588 | 6.4317 | 6.4536 | 6.115 | 6.2017 | 6.6183 | 6.2517 | 6.375 | 6.3067 | 6.185 | 6.5485 | 6.1167 | 6.3992 | 4.045 | 6.0187 | 6.4447 | 6.5617 | 6.6485 | 6.18 | 5.8517 | 6.14 | 6.8667
| Median |  | 6.18 | 6.5922 | 6.4237 | 6.3144 | 6.03 | 7.147 | 6.415 | 6.4722 | 6.115 | 6.21 | 6.645 | 6.23 | 6.31 | 6.305 | 6.19 | 6.5535 | 6.135 | 6.362 | 4.17 | 5.9944 | 6.4565 | 6.555 | 6.6627 | 6.19 | 5.85 | 6.14 | 6.9
| Std Dev. |  | 0.083586 | 0.07559 | 0.036817 | 0.1561 | 0.18681 | 0.028986 | 0.081097 | 0.11563 | 0.010488 | 0.06047 | 0.062102 | 0.098675 | 0.1878 | 0.054283 | 0.10095 | 0.044649 | 0.037238 | 0.17396 | 0.50386 | 0.11589 | 0.048111 | 0.052694 | 0.087778 | 0.070427 | 0.048339 | 0.038987 | 0.05164
| Rel.Std.Dev. |  | 1.361% | 1.152% | 0.5726% | 2.482% | 3.087% | 0.4049% | 1.261% | 1.792% | 0.1715% | 0.9751% | 0.9383% | 1.578% | 2.946% | 0.8607% | 1.632% | 0.6818% | 0.6088% | 2.718% | 12.46% | 1.925% | 0.7465% | 0.8031% | 1.32% | 1.14% | 0.8261% | 0.635% | 0.752%
| PDM3 |  | -2.874% | 3.768% | 1.653% | -0.5642% | -4.323% | 13.18% | 1.685% | 2.032% | -3.322% | -1.951% | 4.636% | -1.161% | 0.789% | -0.2913% | -2.215% | 3.532% | -3.295% | 1.171% | -36.05% | -4.844% | 1.89% | 3.74% | 5.113% | -2.294% | -7.485% | -2.926% | 8.562%
| Z-Score (Absolute) |  | 0.66 | 1.04 | 0.5 | 0.07 | 1.03 | 3.45 | 0.51 | 0.59 | 0.78 | 0.42 | 1.26 | 0.22 | 0.28 | 0 | 0.49 | 0.98 | 0.77 | 0.37 | 9.15 | 1.17 | 0.56 | 1.03 | 1.38 | 0.51 | 1.84 | 0.67 | 2.27
| Analytical results for Cd in OREAS 233b (Certified Value 0.58 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.67 | 0.59 | 0.25 | 0.58934 | 0.5 | 0.7 | 0.65 | 0.61 | 0.9 | 0.7 | 0.6 | 0.56 | 0.55 | 0.54 | < 1 | 0.55 | 0.5 | 0.53 | < 0.1 | 0.5 | 0.51559 | 0.58 | 0.6789 | 0.56 | 0.67 | 0.65 | 0.56 |  |  | 1
| 1 | 2 | 0.6 | 0.61 | 0.23 | 0.63291 | 0.5 | 0.7 | 0.67 | 0.58 | 0.9 | 0.6 | 0.57 | 0.54 | 0.55 | 0.58 | < 1 | 0.58 | 0.5 | 0.56 | < 0.1 | 0.5 | 0.55133 | 0.54 | 0.6105 | 0.57 | 0.63 | 0.63 | 0.49 |  |  | 23
| 1 | 3 | 0.66 | 0.61 | 0.26 | 0.63067 | 0.4 | 0.7 | 0.61 | 0.6 | 0.9 | 0.6 | 0.6 | 0.53 | 0.55 | 0.61 | < 1 | 0.57 | 0.5 | 0.55 | < 0.1 | 0.5 | 0.54338 | 0.54 | 0.6351 | 0.55 | 0.61 | 0.63 | 0.47 |  |  | 16
| 1 | 4 | 0.64 | 0.56 | 0.27 | 0.62437 | 0.4 | 0.7 | 0.71 | 0.54 | 0.9 | 0.7 | 0.58 | 0.58 | 0.55 | 0.57 | < 1 | 0.49 | 0.5 | 0.55 | < 0.1 | 0.5 | 0.56076 | 0.57 | 0.5931 | 0.59 | 0.63 | 0.66 | 0.47 |  |  | 0.58174
| 1 | 5 | 0.63 | 0.59 | 0.27 | 0.63909 | 0.4 | 0.7 | 0.55 | 0.53 | 0.9 | 0.6 | 0.61 | 0.55 | 0.53 | 0.62 | < 1 | 0.5 | 0.5 | 0.55 | < 0.1 | 0.5 | 0.52983 | 0.57 | 0.577 | 0.56 | 0.59 | 0.65 | 0.49 |  |  | 22
| 1 | 6 | 0.69 | 0.61 | 0.27 | 0.61732 | 0.5 | 0.7 | 0.67 | 0.53 | 0.9 | 0.7 | 0.57 | 0.58 | 0.52 | 0.59 | < 1 | 0.55 | 0.5 | 0.55 | < 0.1 | 0.5 | 0.56337 | 0.56 | 0.6003 | 0.56 | 0.62 | 0.63 | 0.48
| Mean |  | 0.64833 | 0.595 | 0.25833 | 0.62228 | 0.45 | 0.7 | 0.64333 | 0.565 | 0.9 | 0.65 | 0.58833 | 0.55667 | 0.54167 | 0.585 |  | 0.54 | 0.5 | 0.54833 |  | 0.5 | 0.54404 | 0.56 | 0.61582 | 0.565 | 0.625 | 0.64167 | 0.49333
| Median |  | 0.65 | 0.6 | 0.265 | 0.62752 | 0.45 | 0.7 | 0.66 | 0.56 | 0.9 | 0.65 | 0.59 | 0.555 | 0.55 | 0.585 |  | 0.55 | 0.5 | 0.55 |  | 0.5 | 0.54736 | 0.565 | 0.6054 | 0.56 | 0.625 | 0.64 | 0.485
| Std Dev. |  | 0.031885 | 0.019748 | 0.016021 | 0.017778 | 0.054772 | 1.2162e-16 | 0.056095 | 0.036194 | 0 | 0.054772 | 0.017224 | 0.020656 | 0.013292 | 0.02881 |  | 0.036878 | 0 | 0.0098319 |  | 0 | 0.018537 | 0.016733 | 0.03645 | 0.013784 | 0.026646 | 0.013292 | 0.033862
| Rel.Std.Dev. |  | 4.918% | 3.319% | 6.202% | 2.857% | 12.17% | 1.737e-14% | 8.719% | 6.406% | 0% | 8.427% | 2.928% | 3.711% | 2.454% | 4.925% |  | 6.829% | 0% | 1.793% |  | 0% | 3.407% | 2.988% | 5.919% | 2.44% | 4.263% | 2.071% | 6.864%
| PDM3 |  | 11.45% | 2.279% | -55.59% | 6.968% | -22.65% | 20.33% | 10.59% | -2.878% | 54.71% | 11.73% | 1.133% | -4.311% | -6.889% | 0.5596% |  | -7.176% | -14.05% | -5.743% |  | -14.05% | -6.481% | -3.738% | 5.857% | -2.878% | 7.436% | 10.3% | -15.2%
| Z-Score (Absolute) |  | 1.87 | 0.67 | 6.89 | 1.29 | NA | NA | 1.76 | 0 | NA | NA | 0.52 | 0.19 | 0.52 | 0.45 | 1.46 | 0.56 | NA | 0.37 | 11.58 | NA | 0.47 | 0.11 | 1.14 | 0 | 1.35 | 1.72 | 1.61
Results from laboratories 1.07, 1.08, 1.12, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 233b (Certified Value 12.6 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 13.06 | 12.95 | 10.86 | 11.9285 | 12.7788 | 12.5 | 15 | 12.34 | 13 | 12.7 | 13.2 | 11.75 | 11.75 | < 10 | 12 | 12.37 | 14.5 | 12.6578 | 12.9 | 7.2641 | 14.7 | 13.28 | 14.35 | 12.1 | [5 empty] | 1
| 1 | 2 | 14.09 | 12.86 | 11.83 | 11.9706 | 12.8141 | 12.7 | 16 | 12.23 | 13 | 13.2 | 12.8 | 12.65 | 12.35 | < 10 | 12 | 12.29 | 14.5 | 12.6387 | 13.25 | 7.4135 | 14.8 | 13.04 | 14.07 | 11.3 | [5 empty] | 24
| 1 | 3 | 13.39 | 12.8 | 11.62 | 11.8533 | 12.4683 | 12.1 | 16 | 12.25 | 13 | 13.1 | 12.75 | 11.85 | 12.85 | < 10 | 12 | 12.43 | 14.5 | 12.5466 | 12.95 | 7.4067 | 14.7 | 12.63 | 14.23 | 10.9 | [5 empty] | 16
| 1 | 4 | 12.66 | 12.83 | 12.47 | 12.0622 | 12.9325 | 12 | 15 | 12.56 | 13 | 12.85 | 13.45 | 12.25 | 12.85 | < 10 | 12 | 12.44 | 14 | 12.6945 | 12.9 | 7.4245 | 15 | 12.83 | 14.8 | 10.7 | [5 empty] | 12.5654
| 1 | 5 | 12.17 | 12.84 | 11.89 | 11.8934 | 13.1909 | 11.9 | 15 | 12.18 | 13 | 12.9 | 13.4 | 12.55 | 12.8 | < 10 | 12 | 11.92 | 14 | 12.2707 | 13.05 | 7.1736 | 14.9 | 13 | 14.67 | 11.2 | [5 empty] | 23
| 1 | 6 | 12.25 | 12.9 | 12.27 | 11.9218 | 12.4244 | 11.9 | 15 | 12.47 | 13 | 12.65 | 13.05 | 12.1 | 12.55 | < 10 | 12 | 12 | 13.5 | 12.8078 | 12.8 | 7.1666 | 14.8 | 12.82 | 14.8 | 10.5
| Mean |  | 12.9367 | 12.8633 | 11.8233 | 11.9383 | 12.7682 | 12.1833 | 15.3333 | 12.3383 | 13 | 12.9 | 13.1083 | 12.1917 | 12.525 |  | 12 | 12.2417 | 14.1667 | 12.6027 | 12.975 | 7.3082 | 14.8167 | 12.9333 | 14.4867 | 11.1167
| Median |  | 12.86 | 12.85 | 11.86 | 11.9251 | 12.7965 | 12.05 | 15 | 12.295 | 13 | 12.875 | 13.125 | 12.175 | 12.675 |  | 12 | 12.33 | 14.25 | 12.6482 | 12.925 | 7.3354 | 14.8 | 12.915 | 14.51 | 11.05
| Std Dev. |  | 0.73304 | 0.053914 | 0.56419 | 0.072089 | 0.28844 | 0.33714 | 0.5164 | 0.14905 | 0 | 0.21679 | 0.29566 | 0.36389 | 0.42866 |  | 0 | 0.22605 | 0.40825 | 0.18342 | 0.15732 | 0.12201 | 0.1169 | 0.22429 | 0.31245 | 0.56716
| Rel.Std.Dev. |  | 5.666% | 0.4191% | 4.772% | 0.6038% | 2.259% | 2.767% | 3.368% | 1.208% | 0% | 1.681% | 2.256% | 2.985% | 3.422% |  | 0% | 1.847% | 2.882% | 1.455% | 1.212% | 1.669% | 0.789% | 1.734% | 2.157% | 5.102%
| PDM3 |  | 2.954% | 2.371% | -5.906% | -4.991% | 1.613% | -3.041% | 22.03% | -1.807% | 3.458% | 2.663% | 4.321% | -2.975% | -0.3218% |  | -4.5% | -2.577% | 12.74% | 0.2963% | 3.259% | -41.84% | 17.92% | 2.928% | 15.29% | -11.53%
| Z-Score (Absolute) |  | 0.55 | 0.43 | 1.28 | 1.09 | 0.27 | 0.69 | NA | 0.43 | NA | 0.49 | 0.83 | 0.67 | 0.13 | 12.47 | NA | 0.59 | 2.57 | 0 | 0.61 | 8.69 | 3.63 | 0.54 | 3.09 | 2.44
Results from laboratories 1.08, 1.13 and 1.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 233b (Certified Value 43.6 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 47.3 | 45.1 | 40.64 | 36.5138 | 42.588 | 43 | 54.6 | 45 | 44.1 | 34 | 43.3 | 44 | 41.9 | 42.9 | 40.7 | 42 | 43.9 | 41.1 | 45.3 | 36.88 | 40 | 43.5144 | 43.1 | 44.2469 | 49.9 | 46.4 | 50.4 | 41.3 |  | 1
| 1 | 2 | 46.2 | 45.1 | 39.95 | 37.3152 | 43.3467 | 43 | 51.3 | 44 | 44.6 | 33 | 43.1 | 45.4 | 40.6 | 45.6 | 43.5 | 43 | 41.7 | 41.7 | 46 | 35.95 | 40 | 45.1857 | 43.8 | 43.2487 | 50.7 | 45.9 | 49.1 | 39.2 |  | 25
| 1 | 3 | 46.8 | 44.5 | 42.47 | 36.5886 | 42.8509 | 43 | 52 | 45 | 44 | 33 | 43.5 | 45.2 | 44.2 | 44.1 | 43.5 | 43 | 42.2 | 41.4 | 46.1 | 37.16 | 40 | 45.6068 | 43.8 | 43.8304 | 50.6 | 44.9 | 49.7 | 37.7 |  | 16
| 1 | 4 | 46 | 44.5 | 52.01 | 37.0153 | 43.0781 | 42 | 50.7 | 44 | 44 | 33 | 43.4 | 43.8 | 41.1 | 42.3 | 44.3 | 43 | 42.6 | 41.2 | 44.3 | 36.4 | 40 | 43.9155 | 43.6 | 43.2606 | 51.9 | 45.4 | 51.6 | 37 |  | 43.5587
| 1 | 5 | 45.8 | 43.7 | 41.36 | 36.9782 | 44.3812 | 43 | 50.8 | 44 | 43.8 | 33 | 43.9 | 44.8 | 42.7 | 44.1 | 45.3 | 44 | 43.2 | 41.8 | 44.1 | 37.96 | 40 | 43.0227 | 45.1 | 42.1557 | 51.3 | 45.5 | 51.4 | 38.2 |  | 24
| 1 | 6 | 47.7 | 44.6 | 42.21 | 35.9328 | 40.8246 | 42 | 49.9 | 45 | 44.6 | 33 | 43.7 | 44.7 | 46.8 | 43.8 | 44.5 | 42 | 43 | 42.3 | 45.9 | 36.85 | 40 | 44.0979 | 44.4 | 44.1855 | 50.5 | 44.3 | 51.2 | 36.4
| Mean |  | 46.6333 | 44.5833 | 43.1067 | 36.724 | 42.8449 | 42.6667 | 51.55 | 44.5 | 44.1833 | 33.1667 | 43.4833 | 44.65 | 42.8833 | 43.8 | 43.6333 | 42.8333 | 42.7667 | 41.5833 | 45.2833 | 36.8667 | 40 | 44.2238 | 43.9667 | 43.488 | 50.8167 | 45.4 | 50.5667 | 38.3
| Median |  | 46.5 | 44.55 | 41.785 | 36.7834 | 42.9645 | 43 | 51.05 | 44.5 | 44.05 | 33 | 43.45 | 44.75 | 42.3 | 43.95 | 43.9 | 43 | 42.8 | 41.55 | 45.6 | 36.865 | 40 | 44.0067 | 43.8 | 43.5455 | 50.65 | 45.45 | 50.8 | 37.95
| Std Dev. |  | 0.7607 | 0.51543 | 4.4627 | 0.48735 | 1.1674 | 0.5164 | 1.6477 | 0.54772 | 0.33714 | 0.40825 | 0.28577 | 0.63797 | 2.3008 | 1.1384 | 1.5883 | 0.75277 | 0.77632 | 0.4446 | 0.88638 | 0.6841 | 0 | 0.98941 | 0.69474 | 0.78277 | 0.69402 | 0.73756 | 1.0093 | 1.7595
| Rel.Std.Dev. |  | 1.631% | 1.156% | 10.35% | 1.327% | 2.725% | 1.21% | 3.196% | 1.231% | 0.7631% | 1.231% | 0.6572% | 1.429% | 5.365% | 2.599% | 3.64% | 1.757% | 1.815% | 1.069% | 1.957% | 1.856% | 0% | 2.237% | 1.58% | 1.8% | 1.366% | 1.625% | 1.996% | 4.594%
| PDM3 |  | 7.059% | 2.352% | -1.038% | -15.69% | -1.639% | -2.048% | 18.35% | 2.161% | 1.434% | -23.86% | -0.173% | 2.505% | -1.55% | 0.554% | 0.1713% | -1.665% | -1.818% | -4.535% | 3.959% | -15.36% | -8.17% | 1.527% | 0.9366% | -0.1624% | 16.66% | 4.227% | 16.09% | -12.07%
| Z-Score (Absolute) |  | 2.11 | 0.7 | 0.31 | 4.7 | 0.49 | 0.61 | 5.49 | 0.65 | 0.43 | 7.14 | 0.05 | 0.75 | 0.47 | 0.16 | 0.05 | 0.5 | 0.55 | 1.36 | 1.18 | 4.6 | 2.45 | 0.46 | 0.28 | 0.05 | 4.99 | 1.26 | 4.82 | 3.62
| Analytical results for Cr in OREAS 233b (Certified Value 146 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 121 | 165 | 135.8637 | 107 | 215 | 163 | 127 | 156 | 150 | 142 | 142 | 139 | 130 | 147 | 123 | 142 | 152 | 140.97 | 137.628 | 165 | 174.852 | 142 | 149.49 | 139 | 162 | 157 | 136 |  |  | 1
| 1 | 2 | 111 | 174 | 136.5759 | 104 | 215 | 156 | 132 | 155 | 151 | 143 | 141 | 133 | 131 | 153 | 125 | 143 | 148 | 136.98 | 143.217 | 160 | 178.7179 | 143 | 153.3876 | 142 | 158 | 164 | 134 |  |  | 26
| 1 | 3 | 112 | 179 | 135.7413 | 104 | 212 | 156 | 137 | 156 | 148 | 142 | 139 | 137 | 132 | 155 | 128 | 144 | 148 | 138.58 | 161.226 | 160 | 178.8007 | 143 | 153.1858 | 158 | 159 | 165 | 136 |  |  | 16
| 1 | 4 | 113 | 180 | 136.893 | 107 | 209 | 142 | 130 | 155 | 153 | 145 | 141 | 139 | 134 | 148 | 128 | 146 | 146 | 137.75 | 165.78 | 160 | 175.9985 | 145 | 152.59 | 151 | 157 | 163 | 135 |  |  | 146.0186
| 1 | 5 | 111 | 173 | 138.6789 | 141 | 213 | 141 | 147 | 153 | 147 | 141 | 142 | 139 | 131 | 156 | 131 | 145 | 173 | 128.48 | 156.699 | 160 | 172.5938 | 142 | 149.7883 | 142 | 153 | 160 | 138 |  |  | 25
| 1 | 6 | 112 | 171 | 128.5254 | 91 | 215 | 147 | 116 | 157 | 152 | 143 | 149 | 150 | 130 | 150 | 124 | 146 | 139 | 123.77 | 165.132 | 155 | 178.3766 | 145 | 151.6875 | 140 | 157 | 158 | 140
| Mean |  | 113.3333 | 173.6667 | 135.3797 | 109 | 213.1667 | 150.8333 | 131.5 | 155.3333 | 150.1667 | 142.6667 | 142.3333 | 139.5 | 131.3333 | 151.5 | 126.5 | 144.3333 | 151 | 134.4217 | 154.947 | 160 | 176.5566 | 143.3333 | 151.6882 | 145.3333 | 157.6667 | 161.1667 | 136.5
| Median |  | 112 | 173.5 | 136.2198 | 105.5 | 214 | 151.5 | 131 | 155.5 | 150.5 | 142.5 | 141.5 | 139 | 131 | 151.5 | 126.5 | 144.5 | 148 | 137.365 | 158.9625 | 160 | 177.1875 | 143 | 152.1388 | 142 | 157.5 | 161.5 | 136
| Std Dev. |  | 3.8297 | 5.5015 | 3.52 | 16.769 | 2.4014 | 8.8412 | 10.3296 | 1.3663 | 2.3166 | 1.3663 | 3.4448 | 5.648 | 1.5055 | 3.7283 | 3.0166 | 1.633 | 11.5931 | 6.7314 | 11.8417 | 3.1623 | 2.5275 | 1.3663 | 1.6962 | 7.5277 | 2.9439 | 3.3116 | 2.1679
| Rel.Std.Dev. |  | 3.379% | 3.168% | 2.6% | 15.38% | 1.127% | 5.862% | 7.855% | 0.8796% | 1.543% | 0.9577% | 2.42% | 4.049% | 1.146% | 2.461% | 2.385% | 1.131% | 7.678% | 5.008% | 7.642% | 1.976% | 1.432% | 0.9532% | 1.118% | 5.18% | 1.867% | 2.055% | 1.588%
| PDM3 |  | -22.38% | 18.93% | -7.286% | -25.35% | 45.99% | 3.297% | -9.943% | 6.379% | 2.841% | -2.296% | -2.524% | -4.464% | -10.06% | 3.754% | -13.37% | -1.154% | 3.412% | -7.942% | 6.115% | 9.575% | 20.91% | -1.839% | 3.883% | -0.4693% | 7.977% | 10.37% | -6.519%
| Z-Score (Absolute) |  | 2.17 | 1.92 | 0.67 | 2.46 | 4.6 | 0.37 | 0.94 | 0.68 | 0.33 | 0.18 | 0.2 | 0.4 | 0.95 | 0.42 | 1.28 | 0.07 | 0.38 | 0.74 | 0.65 | 0.99 | 2.12 | 0.14 | 0.43 | 0 | 0.84 | 1.07 | 0.6
| Analytical results for Cs in OREAS 233b (Certified Value 0.89 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | < 5 | 0.9 | 0.80959 | 0.9 | 1 | 0.9 | 0.87 | 1 | 0.86 | 0.93 | 0.87 | 0.84 | 0.9 | 0.85 | 1 | 0.81405 | 0.92 | 0.9414 | 0.84 | 1.18 | 0.97 | 0.79 | [7 empty] | 1
| 1 | 2 | < 5 | 0.87 | 0.84029 | 0.88 | 0.9 | 0.9 | 0.92 | 1 | 0.9 | 0.92 | 0.9 | 0.84 | 0.8 | 0.88 | 1 | 0.78887 | 0.91 | 0.9651 | 0.83 | 1.06 | 0.98 | 0.74 | [7 empty] | 27
| 1 | 3 | < 5 | 0.89 | 0.84045 | 0.86 | 1 | 0.9 | 0.91 | 0.9 | 0.88 | 0.88 | 0.79 | 0.89 | 0.9 | 0.87 | 1 | 0.78282 | 0.92 | 0.9693 | 0.83 | 1.04 | 0.94 | 0.71 | [7 empty] | 16
| 1 | 4 | < 5 | 0.87 | 0.86844 | 0.83 | 1 | 0.9 | 0.92 | 1 | 0.87 | 0.95 | 0.88 | 0.92 | 0.8 | 0.87 | 1 | 0.81918 | 0.89 | 0.9399 | 0.82 | 1.03 | 1 | 0.72 | [7 empty] | 0.89133
| 1 | 5 | < 5 | 0.88 | 0.79456 | 0.79 | 1.1 | 0.9 | 0.84 | 1 | 0.87 | 0.95 | 0.89 | 0.91 | 0.9 | 0.83 | 1 | 0.78662 | 0.93 | 0.9359 | 0.82 | 1.08 | 1 | 0.74 | [7 empty] | 26
| 1 | 6 | < 5 | 0.9 | 0.816 | 0.81 | 0.9 | 0.8 | 0.89 | 1.1 | 0.88 | 0.89 | 0.87 | 0.92 | 0.9 | 0.86 | 1 | 0.82911 | 0.91 | 0.9401 | 0.83 | 1.01 | 1.02 | 0.71
| Mean |  |  | 0.885 | 0.82822 | 0.845 | 0.98333 | 0.88333 | 0.89167 | 1 | 0.87667 | 0.92 | 0.86667 | 0.88667 | 0.86667 | 0.86 | 1 | 0.80344 | 0.91333 | 0.94862 | 0.82833 | 1.0667 | 0.985 | 0.735
| Median |  |  | 0.885 | 0.82814 | 0.845 | 1 | 0.9 | 0.9 | 1 | 0.875 | 0.925 | 0.875 | 0.9 | 0.9 | 0.865 | 1 | 0.80146 | 0.915 | 0.94075 | 0.83 | 1.05 | 0.99 | 0.73
| Std Dev. |  |  | 0.013784 | 0.026619 | 0.042308 | 0.075277 | 0.040825 | 0.031885 | 0.063246 | 0.013663 | 0.029665 | 0.039328 | 0.037771 | 0.05164 | 0.017889 | 0 | 0.019696 | 0.013663 | 0.014573 | 0.0075277 | 0.060553 | 0.028107 | 0.030166
| Rel.Std.Dev. |  |  | 1.558% | 3.214% | 5.007% | 7.655% | 4.622% | 3.576% | 6.325% | 1.558% | 3.224% | 4.538% | 4.26% | 5.958% | 2.08% | 0% | 2.451% | 1.496% | 1.536% | 0.9088% | 5.677% | 2.853% | 4.104%
| PDM3 |  |  | -0.7102% | -7.08% | -5.198% | 10.32% | -0.8972% | 0.03773% | 12.19% | -1.645% | 3.217% | -2.767% | -0.5232% | -2.767% | -3.515% | 12.19% | -9.86% | 2.469% | 6.427% | -7.068% | 19.67% | 10.51% | -17.54%
| Z-Score (Absolute) |  | 27.23 | 0 | 0.96 | 0.67 | 1.66 | 0.03 | 0.11 | 1.94 | 0.14 | 0.59 | 0.31 | 0.03 | 0.31 | 0.42 | NA | 1.37 | 0.48 | 1.07 | 0.96 | 3.06 | 1.69 | 2.53
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 233b (Certified Value 161 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 167.9 | 160.3 | 152.38 | 171.2 | 159.2578 | 164 | 185 | 156 | 157.8 | 162 | 159.5 | 166.5 | 160.5 | 157.5 | 162 | 165 | 160 | 144.3 | 165 | 20.01 | 150.552 | 155 | 147.9113 | 161 | 151.4997 | 175 | 166.3 | 160.8 | 158 | 1
| 1 | 2 | 166.1 | 160.7 | 141.72 | 173.54 | 160.4057 | 171 | 180 | 153 | 158.2 | 161 | 162.6 | 170 | 157.5 | 164 | 164.5 | 168 | 161 | 146.4 | 169 | 20.8 | 150.174 | 155 | 156.9581 | 157 | 136.3007 | 174 | 163.2 | 160.8 | 159 | 28
| 1 | 3 | 167.4 | 161.4 | 151.88 | 170.43 | 157.7473 | 164 | 183 | 154 | 158.1 | 162 | 156.9 | 172 | 157.5 | 160.5 | 164 | 166 | 161 | 147 | 162 | 20.91 | 158.58 | 155 | 157.7929 | 159.5 | 137.7635 | 168 | 161.1 | 160.6 | 155 | 16
| 1 | 4 | 165.5 | 160.7 | 149.25 | 172.58 | 160.5717 | 164 | 180 | 156 | 162.7 | 162 | 166 | 167.5 | 157 | 157.5 | 165.5 | 168 | 162 | 145.2 | 162 | 19.6 | 157.068 | 155 | 154.8442 | 157 | 134.4658 | 175 | 160.6 | 161.5 | 153 | 160.8734
| 1 | 5 | 162.5 | 160.3 | 150.57 | 172.19 | 162.7012 | 177 | 174 | 156 | 154.4 | 162 | 169.6 | 171.5 | 161 | 160.5 | 165 | 172 | 160 | 145.7 | 173 | 27.07 | 152.622 | 155 | 146.812 | 160.5 | 130.7941 | 174 | 158.2 | 161.6 | 158 | 27
| 1 | 6 | 165.2 | 160.8 | 150.77 | 170.13 | 151.7641 | 165 | 175 | 158 | 159.3 | 163 | 161.3 | 170.5 | 163.5 | 165.5 | 167 | 168 | 162 | 143.4 | 163 | 23.54 | 158.481 | 160 | 150.0432 | 160 | 139.1639 | 177 | 156.7 | 161 | 158
| Mean |  | 165.7667 | 160.7 | 149.4283 | 171.6783 | 158.7413 | 167.5 | 179.5 | 155.5 | 158.4167 | 162 | 162.65 | 169.6667 | 159.5 | 160.9167 | 164.6667 | 167.8333 | 161 | 145.3333 | 165.6667 | 21.9883 | 154.5795 | 155.8333 | 152.3936 | 159.1667 | 138.3313 | 173.8333 | 161.0167 | 161.05 | 156.8333
| Median |  | 165.8 | 160.7 | 150.67 | 171.695 | 159.8317 | 164.5 | 180 | 156 | 158.15 | 162 | 161.95 | 170.25 | 159 | 160.5 | 164.75 | 168 | 161 | 145.45 | 164 | 20.855 | 154.845 | 155 | 152.4437 | 159.75 | 137.0321 | 174.5 | 160.85 | 160.9 | 158
| Std Dev. |  | 1.918 | 0.40497 | 3.9313 | 1.321 | 3.7878 | 5.3944 | 4.3243 | 1.7607 | 2.6754 | 0.63246 | 4.568 | 2.2061 | 2.5884 | 3.2927 | 1.6633 | 2.4014 | 0.89443 | 1.3322 | 4.4572 | 2.8434 | 3.9212 | 2.0412 | 4.7489 | 1.7512 | 7.0753 | 3.0605 | 3.4476 | 0.40866 | 2.3166
| Rel.Std.Dev. |  | 1.157% | 0.252% | 2.631% | 0.7695% | 2.386% | 3.221% | 2.409% | 1.132% | 1.689% | 0.3904% | 2.809% | 1.3% | 1.623% | 2.046% | 1.01% | 1.431% | 0.5555% | 0.9166% | 2.69% | 12.93% | 2.537% | 1.31% | 3.116% | 1.1% | 5.115% | 1.761% | 2.141% | 0.2537% | 1.477%
| PDM3 |  | 3.042% | -0.1078% | -7.114% | 6.716% | -1.325% | 4.119% | 11.58% | -3.34% | -1.527% | 0.7003% | 1.104% | 5.466% | -0.8537% | 0.0269% | 2.358% | 4.326% | 0.0787% | -9.66% | 2.98% | -86.33% | -3.912% | -3.133% | -5.271% | -1.061% | -14.01% | 8.056% | 0.08906% | 0.1098% | -2.511%
| Z-Score (Absolute) |  | 0.67 | 0.03 | 1.6 | 1.5 | 0.3 | 0.92 | 2.58 | 0.75 | 0.35 | 0.15 | 0.24 | 1.22 | 0.2 | 0 | 0.52 | 0.96 | 0.01 | 2.17 | 0.66 | 19.32 | 0.88 | 0.71 | 1.18 | 0.24 | 3.14 | 1.8 | 0.01 | 0.02 | 0.57
| Analytical results for Dy in OREAS 233b (Certified Value 3.6 ppm) | [30 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*OES; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS | 2
3
| 1 | 1 | 1.01: 3.67; 1.02: 3.02; 1.03: 3.5295; 1.04: 3.5; 1.05: 3.91; 1.07: 3.69; 1.08: 3.9; 1.10: 3.7103 | 1
| 1 | 2 | 1.01: 3.67; 1.02: 3.39; 1.03: 3.4771; 1.04: 3.5; 1.05: 3.72; 1.07: 3.58; 1.08: 3.6; 1.10: 3.6388 | 29
| 1 | 3 | 1.01: 3.6; 1.02: 3.32; 1.03: 3.4708; 1.04: 3.6; 1.05: 3.76; 1.07: 3.59; 1.08: 3.8; 1.10: 3.6048 | 16
| 1 | 4 | 1.01: 3.71; 1.02: 3.51; 1.03: 3.4455; 1.04: 3.5; 1.05: 3.76; 1.07: 3.7; 1.08: 3.9; 1.10: 3.6263 | 3.6044
| 1 | 5 | 1.01: 3.67; 1.02: 3.37; 1.03: 3.4571; 1.04: 3.6; 1.05: 3.71; 1.07: 3.57; 1.08: 3.8; 1.10: 3.5342 | 28
| 1 | 6 | 1.01: 3.69; 1.02: 3.16; 1.03: 3.5025; 1.04: 3.6; 1.05: 3.6; 1.07: 3.54; 1.08: 3.7; 1.10: 3.774
| Mean |  | 1.01: 3.6683; 1.02: 3.295; 1.03: 3.4804; 1.04: 3.55; 1.05: 3.7433; 1.07: 3.6117; 1.08: 3.7833; 1.10: 3.6481
| Median |  | 1.01: 3.67; 1.02: 3.345; 1.03: 3.4739; 1.04: 3.55; 1.05: 3.74; 1.07: 3.585; 1.08: 3.8; 1.10: 3.6326
| Std Dev. |  | 1.01: 0.037103; 1.02: 0.17627; 1.03: 0.030864; 1.04: 0.054772; 1.05: 0.10053; 1.07: 0.066758; 1.08: 0.1169; 1.10: 0.083845
| Rel.Std.Dev. |  | 1.01: 1.011%; 1.02: 5.35%; 1.03: 0.8868%; 1.04: 1.543%; 1.05: 2.686%; 1.07: 1.848%; 1.08: 3.09%; 1.10: 2.298%
| PDM3 |  | 1.01: 1.774%; 1.02: -8.584%; 1.03: -3.44%; 1.04: -1.509%; 1.05: 3.855%; 1.07: 0.2018%; 1.08: 4.965%; 1.10: 1.212%
| Z-Score (Absolute) |  | 1.01: 0.27; 1.02: 2.34; 1.03: 1.04; 1.04: 0.56; 1.05: 0.79; 1.07: 0.13; 1.08: 1.07; 1.10: 0.13
| Analytical results for Er in OREAS 233b (Certified Value 2.18 ppm) | [30 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*OES; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS | 2
3
| 1 | 1 | 1.01: 2.2; 1.02: 1.8; 1.03: 2.1146; 1.04: 2.1; 1.05: 2.4; 1.07: 2.17; 1.08: 2.4; 1.10: 2.3296 | 1
| 1 | 2 | 1.01: 2.25; 1.02: 2.02; 1.03: 2.114; 1.04: 2.1; 1.05: 2.27; 1.07: 2.14; 1.08: 2.3; 1.10: 2.3097 | 30
| 1 | 3 | 1.01: 2.18; 1.02: 1.97; 1.03: 2.1103; 1.04: 2.2; 1.05: 2.3; 1.07: 2.19; 1.08: 2.3; 1.10: 2.2696 | 16
| 1 | 4 | 1.01: 2.26; 1.02: 2.08; 1.03: 2.1457; 1.04: 2.1; 1.05: 2.29; 1.07: 2.17; 1.08: 2.2; 1.10: 2.3364 | 2.1847
| 1 | 5 | 1.01: 2.17; 1.02: 1.97; 1.03: 2.1495; 1.04: 2.1; 1.05: 2.25; 1.07: 2.14; 1.08: 2.3; 1.10: 2.23 | 29
| 1 | 6 | 1.01: 2.29; 1.02: 1.85; 1.03: 2.1368; 1.04: 2.1; 1.05: 2.18; 1.07: 2.15; 1.08: 2.3; 1.10: 2.4307
| Mean |  | 1.01: 2.225; 1.02: 1.9483; 1.03: 2.1285; 1.04: 2.1167; 1.05: 2.2817; 1.07: 2.16; 1.08: 2.3; 1.10: 2.3177
| Median |  | 1.01: 2.225; 1.02: 1.97; 1.03: 2.1257; 1.04: 2.1; 1.05: 2.28; 1.07: 2.16; 1.08: 2.3; 1.10: 2.3196
| Std Dev. |  | 1.01: 0.048477; 1.02: 0.10496; 1.03: 0.017533; 1.04: 0.040825; 1.05: 0.071949; 1.07: 0.02; 1.08: 0.063246; 1.10: 0.068309
| Rel.Std.Dev. |  | 1.01: 2.179%; 1.02: 5.387%; 1.03: 0.8237%; 1.04: 1.929%; 1.05: 3.153%; 1.07: 0.9259%; 1.08: 2.75%; 1.10: 2.947%
| PDM3 |  | 1.01: 1.843%; 1.02: -10.82%; 1.03: -2.574%; 1.04: -3.115%; 1.05: 4.437%; 1.07: -1.132%; 1.08: 5.276%; 1.10: 6.085%
| Z-Score (Absolute) |  | 1.01: 0.27; 1.02: 2; 1.03: 0.52; 1.04: 0.62; 1.05: 0.73; 1.07: 0.27; 1.08: 0.88; 1.10: 1.02
| Analytical results for Eu in OREAS 233b (Certified Value 0.9 ppm) | [30 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*OES; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS | 2
3
| 1 | 1 | 1.01: 0.88; 1.02: 0.77; 1.03: 1.2578; 1.04: 0.85; 1.05: 0.99; 1.07: 0.96; 1.08: 0.9; 1.10: 0.90964 | 1
| 1 | 2 | 1.01: 0.87; 1.02: 0.86; 1.03: 1.2916; 1.04: 0.84; 1.05: 0.95; 1.07: 0.96; 1.08: 0.9; 1.10: 0.88104 | 31
| 1 | 3 | 1.01: 0.89; 1.02: 0.84; 1.03: 1.2589; 1.04: 0.85; 1.05: 0.95; 1.07: 0.98; 1.08: 1; 1.10: 0.88302 | 16
| 1 | 4 | 1.01: 0.9; 1.02: 0.91; 1.03: 1.2314; 1.04: 0.84; 1.05: 0.95; 1.07: 0.95; 1.08: 1; 1.10: 0.88661 | 0.89567
| 1 | 5 | 1.01: 0.89; 1.02: 0.86; 1.03: 1.2419; 1.04: 0.84; 1.05: 0.96; 1.07: 0.97; 1.08: 1; 1.10: 0.86123 | 30
| 1 | 6 | 1.01: 0.91; 1.02: 0.81; 1.03: 1.2533; 1.04: 0.83; 1.05: 0.91; 1.07: 0.94; 1.08: 0.9; 1.10: 0.9127
| Mean |  | 1.01: 0.89; 1.02: 0.84167; 1.03: 1.2558; 1.04: 0.84167; 1.05: 0.95167; 1.07: 0.96; 1.08: 0.95; 1.10: 0.88904
| Median |  | 1.01: 0.89; 1.02: 0.85; 1.03: 1.2555; 1.04: 0.84; 1.05: 0.95; 1.07: 0.96; 1.08: 0.95; 1.10: 0.88482
| Std Dev. |  | 1.01: 0.014142; 1.02: 0.047924; 1.03: 0.020426; 1.04: 0.0075277; 1.05: 0.025626; 1.07: 0.014142; 1.08: 0.054772; 1.10: 0.019306
| Rel.Std.Dev. |  | 1.01: 1.589%; 1.02: 5.694%; 1.03: 1.626%; 1.04: 0.8944%; 1.05: 2.693%; 1.07: 1.473%; 1.08: 5.766%; 1.10: 2.172%
| PDM3 |  | 1.01: -0.6334%; 1.02: -6.03%; 1.03: 40.21%; 1.04: -6.03%; 1.05: 6.252%; 1.07: 7.182%; 1.08: 6.065%; 1.10: -0.7406%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 5.1; 1.04: 0.67; 1.05: 0.86; 1.07: 0.98; 1.08: NA; 1.10: 0.01
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 233b (Certified Value 7.84 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 8.2 | 8.17 | 8.0182 | 7.7339 | 8.04 | 8.813 | 8.21 | 7.78 | 8.07 | 7.55 | 8.14 | 8.09 | 7.84 | 8.01 | 7.34 | 8.34 | 7.25 | 7.84 | 7.67 | 7.223 | 8 | 8.08 | 7.6463 | 7.71 | 7.55 | 8.34 | 8.9 |  |  | 1
| 1 | 2 | 8.12 | 8.13 | 8.0314 | 7.8344 | 7.79 | 8.743 | 7.95 | 7.88 | 7.88 | 7.65 | 8.39 | 7.84 | 8.03 | 7.94 | 7.55 | 8.33 | 7.35 | 7.86 | 7.45 | 7.2245 | 7.925 | 8.01 | 7.9253 | 7.84 | 7.32 | 8.45 | 8.9 |  |  | 0
| 1 | 3 | 8.14 | 8.17 | 8.0387 | 7.7535 | 7.61 | 8.743 | 8.08 | 7.76 | 7.98 | 7.46 | 8.41 | 7.78 | 7.6 | 8.02 | 7.44 | 8.21 | 7.34 | 7.68 | 7.49 | 7.5589 | 7.925 | 7.98 | 7.8433 | 7.59 | 7.5 | 8.34 | 8.9 |  |  | 16
| 1 | 4 | 8.04 | 8.07 | 8.0685 | 7.8599 | 7.4 | 8.743 | 8.13 | 7.8 | 7.9 | 7.56 | 8.39 | 7.8 | 7.83 | 8.03 | 7.44 | 8.42 | 7.31 | 7.63 | 7.41 | 7.4154 | 7.96 | 7.95 | 7.8983 | 7.82 | 7.28 | 8.38 | 8.9 |  |  | 7.8431
| 1 | 5 | 7.91 | 7.89 | 8.0906 | 7.8101 | 7.25 | 8.673 | 8.15 | 7.78 | 7.92 | 7.54 | 8.43 | 7.99 | 7.76 | 7.89 | 7.6 | 8.25 | 7.35 | 8.21 | 7.16 | 7.2845 | 8.01 | 7.93 | 7.6953 | 7.72 | 7.24 | 8.38 | 8.9 |  |  | 31
| 1 | 6 | 8.17 | 8.09 | 8.0788 | 7.8341 | 7.25 | 8.743 | 8.31 | 7.83 | 7.93 | 7.46 | 8.39 | 7.8 | 8.34 | 8.06 | 7.37 | 8.42 | 7.29 | 7.72 | 7.07 | 7.5781 | 7.94 | 8.03 | 7.8353 | 7.76 | 7.27 | 8.36 | 9
| Mean |  | 8.0967 | 8.0867 | 8.0544 | 7.8043 | 7.5567 | 8.743 | 8.1383 | 7.805 | 7.9467 | 7.5367 | 8.3583 | 7.8833 | 7.9 | 7.9917 | 7.4567 | 8.3283 | 7.315 | 7.8233 | 7.375 | 7.3807 | 7.96 | 7.9967 | 7.8073 | 7.74 | 7.36 | 8.375 | 8.9167
| Median |  | 8.13 | 8.11 | 8.0536 | 7.8221 | 7.505 | 8.743 | 8.14 | 7.79 | 7.925 | 7.545 | 8.39 | 7.82 | 7.835 | 8.015 | 7.44 | 8.335 | 7.325 | 7.78 | 7.43 | 7.3499 | 7.95 | 7.995 | 7.8393 | 7.74 | 7.3 | 8.37 | 8.9
| Std Dev. |  | 0.10633 | 0.10463 | 0.028942 | 0.049918 | 0.31734 | 0.044272 | 0.12139 | 0.043704 | 0.069186 | 0.071181 | 0.10815 | 0.12691 | 0.25636 | 0.063692 | 0.10093 | 0.08612 | 0.039875 | 0.20963 | 0.22197 | 0.16152 | 0.037283 | 0.055015 | 0.11202 | 0.090111 | 0.1313 | 0.040866 | 0.040825
| Rel.Std.Dev. |  | 1.313% | 1.294% | 0.3593% | 0.6396% | 4.2% | 0.5064% | 1.492% | 0.5599% | 0.8706% | 0.9445% | 1.294% | 1.61% | 3.245% | 0.797% | 1.354% | 1.034% | 0.5451% | 2.68% | 3.01% | 2.188% | 0.4684% | 0.688% | 1.435% | 1.164% | 1.784% | 0.4879% | 0.4578%
| PDM3 |  | 3.233% | 3.106% | 2.694% | -0.494% | -3.652% | 11.47% | 3.765% | -0.4854% | 1.321% | -3.907% | 6.57% | 0.5134% | 0.7259% | 1.895% | -4.927% | 6.187% | -6.733% | -0.2516% | -5.968% | -5.895% | 1.491% | 1.958% | -0.4559% | -1.314% | -6.159% | 6.782% | 13.69%
| Z-Score (Absolute) |  | 0.67 | 0.64 | 0.53 | 0.33 | 1.18 | 2.89 | 0.82 | 0.33 | 0.16 | 1.25 | 1.57 | 0.06 | 0 | 0.31 | 1.52 | 1.47 | 2.01 | 0.26 | 1.8 | 1.78 | 0.21 | 0.33 | 0.32 | 0.55 | 1.85 | 1.63 | 3.49
| Analytical results for Ga in OREAS 233b (Certified Value 15 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 17 | 14.99 | 14.416 | 13.4203 | 14.2 | 15.9 | 15.7 | 15.79 | 19 | 15 | 15.55 | 15.35 | 14.75 | 14.3 | 10 | 15.1 | 15.17 | 12.55 | 12 | 12.5268 | 15.95 | 16.3101 | 18.3 | 16.4 | 15.61 | 13.9 |  |  |  | 1
| 1 | 2 | 16.6 | 15.38 | 14.7326 | 13.6015 | 13.7 | 15 | 15.7 | 15.61 | 18 | 15 | 16.05 | 14.15 | 15.8 | 14.85 | 11 | 14.3 | 15.45 | 12.28 | 12 | 12.0152 | 16.15 | 16.3394 | 18.7 | 15.7 | 15.27 | 13.1 |  |  |  | 32
| 1 | 3 | 16.9 | 15.26 | 14.2597 | 13.5134 | 13.5 | 15.1 | 15.4 | 15.32 | 18 | 15 | 16.05 | 15 | 14.55 | 15.4 | 11 | 14.9 | 15.81 | 12.62 | 12 | 11.597 | 15.8 | 16.5145 | 18.6 | 15.8 | 15.83 | 12.8 |  |  |  | 16
| 1 | 4 | 16.4 | 15.08 | 14.7594 | 13.6383 | 13.3 | 15.1 | 15.6 | 15.7 | 18 | 15 | 15.75 | 15.1 | 14.9 | 15.55 | 11 | 14.5 | 16 | 14.32 | 12 | 12.4661 | 15.6 | 16.4408 | 19.2 | 15.6 | 16.26 | 12.4 |  |  |  | 15.0173
| 1 | 5 | 16.4 | 14.64 | 14.5203 | 13.9598 | 12.9 | 14.9 | 15.2 | 15.28 | 18 | 15 | 15.75 | 14.95 | 15.3 | 15.35 | 11 | 14.5 | 15.35 | 13.06 | 13 | 12.2572 | 16.05 | 16.2006 | 18.9 | 15.9 | 16.29 | 13.1 |  |  |  | 32
| 1 | 6 | 17.2 | 15.49 | 14.3384 | 12.9488 | 12.9 | 14.6 | 15.6 | 15.39 | 18 | 15 | 15.55 | 16.55 | 14.8 | 15.85 | 11 | 14.8 | 15.63 | 9.3 | 12 | 13.0852 | 16.05 | 17.0731 | 18.7 | 15.3 | 16.1 | 12.2
| Mean |  | 16.75 | 15.14 | 14.5044 | 13.5137 | 13.4167 | 15.1 | 15.5333 | 15.515 | 18.1667 | 15 | 15.7833 | 15.1833 | 15.0167 | 15.2167 | 10.8333 | 14.6833 | 15.5683 | 12.355 | 12.1667 | 12.3246 | 15.9333 | 16.4797 | 18.7333 | 15.7833 | 15.8933 | 12.9167
| Median |  | 16.75 | 15.17 | 14.4682 | 13.5574 | 13.4 | 15.05 | 15.6 | 15.5 | 18 | 15 | 15.75 | 15.05 | 14.85 | 15.375 | 11 | 14.65 | 15.54 | 12.585 | 12 | 12.3617 | 16 | 16.3901 | 18.7 | 15.75 | 15.965 | 12.95
| Std Dev. |  | 0.33317 | 0.30679 | 0.20617 | 0.33166 | 0.49967 | 0.43359 | 0.19664 | 0.21342 | 0.40825 | 0 | 0.22509 | 0.78209 | 0.4568 | 0.55468 | 0.40825 | 0.29944 | 0.30623 | 1.6616 | 0.40825 | 0.5038 | 0.20166 | 0.31022 | 0.30111 | 0.3656 | 0.40183 | 0.6047
| Rel.Std.Dev. |  | 1.989% | 2.026% | 1.421% | 2.454% | 3.724% | 2.871% | 1.266% | 1.376% | 2.247% | 0% | 1.426% | 5.151% | 3.042% | 3.645% | 3.768% | 2.039% | 1.967% | 13.45% | 3.355% | 4.088% | 1.266% | 1.882% | 1.607% | 2.316% | 2.528% | 4.682%
| PDM3 |  | 11.54% | 0.8173% | -3.415% | -10.01% | -10.66% | 0.551% | 3.437% | 3.314% | 20.97% | -0.1149% | 5.101% | 1.106% | -0.003939% | 1.328% | -27.86% | -2.224% | 3.67% | -17.73% | -18.98% | -17.93% | 6.1% | 9.739% | 24.75% | 5.101% | 5.834% | -13.99%
| Z-Score (Absolute) |  | 1.64 | 0.06 | 0.73 | 1.78 | 1.88 | 0.11 | 0.35 | 0.33 | NA | NA | 0.62 | 0.02 | 0.19 | 0.02 | NA | 0.55 | 0.39 | 3.01 | NA | 3.04 | 0.77 | 1.35 | 3.73 | 0.62 | 0.73 | 2.41
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 233b (Certified Value 2.96 ppm) | [30 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab; 1.11: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.08; 1.08: 1.24; 1.10: 1.27; 1.11: 1.28 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.07: 4A*OES; 1.08: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.01: 3.21; 1.02: 2.49; 1.04: 3.3153; 1.05: 2.7; 1.07: 2.9; 1.08: 3.13; 1.10: 2.9; 1.11: 0.55859 |  |  |  | 1
| 1 | 2 | 1.01: 3.1; 1.02: 2.78; 1.04: 3.2984; 1.05: 2.6; 1.07: 2.7; 1.08: 3.26; 1.10: 3; 1.11: 0.53237 |  |  |  | 33
| 1 | 3 | 1.01: 3.07; 1.02: 2.73; 1.04: 3.288; 1.05: 2.8; 1.07: 2.8; 1.08: 3.12; 1.10: 2.9; 1.11: 0.54313 |  |  |  | 16
| 1 | 4 | 1.01: 3.1; 1.02: 2.86; 1.04: 3.2996; 1.05: 2.7; 1.07: 2.8; 1.08: 3.17; 1.10: 3.1; 1.11: 0.5453 |  |  |  | 2.9573
| 1 | 5 | 1.01: 3.05; 1.02: 2.76; 1.04: 3.3456; 1.05: 2.7; 1.07: 2.8; 1.08: 3.12; 1.10: 2.9; 1.11: 0.53732 |  |  |  | 33
| 1 | 6 | 1.01: 3.31; 1.02: 2.59; 1.04: 3.292; 1.05: 2.8; 1.07: 2.6; 1.08: 3.12; 1.10: 2.9; 1.11: 0.55182
| Mean |  | 1.01: 3.14; 1.02: 2.7017; 1.04: 3.3065; 1.05: 2.7167; 1.07: 2.7667; 1.08: 3.1533; 1.10: 2.95; 1.11: 0.54475
| Median |  | 1.01: 3.1; 1.02: 2.745; 1.04: 3.299; 1.05: 2.7; 1.07: 2.8; 1.08: 3.125; 1.10: 2.9; 1.11: 0.54422
| Std Dev. |  | 1.01: 0.1; 1.02: 0.13615; 1.04: 0.021311; 1.05: 0.075277; 1.07: 0.10328; 1.08: 0.055737; 1.10: 0.083666; 1.11: 0.0095176
| Rel.Std.Dev. |  | 1.01: 3.185%; 1.02: 5.039%; 1.04: 0.6445%; 1.05: 2.771%; 1.07: 3.733%; 1.08: 1.768%; 1.10: 2.836%; 1.11: 1.747%
| PDM3 |  | 1.01: 6.179%; 1.02: -8.643%; 1.04: 11.81%; 1.05: -8.136%; 1.07: -6.445%; 1.08: 6.63%; 1.10: -0.2454%; 1.11: -81.58%
| Z-Score (Absolute) |  | 1.01: 0.87; 1.02: 0.48; 1.04: 1.38; 1.05: 0.44; 1.07: 0.28; 1.08: 0.91; 1.10: 0.28; 1.11: 7.12
| Analytical results for Ge in OREAS 233b (Certified Value 0.083 ppm) | [30 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | < 0.1 | 1.7 | 0.1416 | < 0.2 | 0.4 | 1.4 | 1 | 3 | 0.09 | 0.09 | 0.08 | 0.07 | < 10 | 0.7 | < 5 | 0.06 | 0.2544 | 0.11 | 0.2 | < 0.05 | 0.09 | [8 empty] | 1
| 1 | 2 | < 0.1 | 1.7 | 0.13193 | < 0.2 | 0.4 | 1.5 | 1 | 2.7 | 0.08 | 0.09 | 0.13 | 0.08 | < 10 | 0.7 | < 5 | 0.07 | 0.2351 | 0.1 | 0.2 | 0.06 | 0.09 | [8 empty] | 34
| 1 | 3 | < 0.1 | 1.8 | 0.13707 | < 0.2 | 0.3 | 1.5 | 1.1 | 3 | 0.08 | 0.12 | 0.07 | 0.07 | < 10 | 0.7 | < 5 | 0.1 | 0.2126 | 0.12 | 0.2 | 0.06 | 0.09 | [8 empty] | 16
| 1 | 4 | < 0.1 | 1.7 | 0.13997 | < 0.2 | 0.2 | 1.6 | 1.3 | 3 | 0.09 | 0.1 | 0.06 | 0.07 | < 10 | 0.6 | < 5 | 0.09 | 0.2098 | 0.12 | 0.2 | < 0.05 | 0.08 | [8 empty] | 0.08275
| 1 | 5 | < 0.1 | 1.7 | 0.12613 | < 0.2 | 0.1 | 1.7 | 1.3 | 2.9 | 0.09 | 0.09 | 0.1 | 0.07 | < 10 | 0.6 | < 5 | 0.06 | 0.1673 | 0.09 | 0.3 | 0.06 | 0.1 | [8 empty] | 34
| 1 | 6 | < 0.1 | 1.7 | 0.13033 | < 0.2 | 0.2 | 1.8 | 1.1 | 2.8 | 0.08 | 0.09 | 0.07 | 0.09 | < 10 | 0.6 | < 5 | 0.06 | 0.1987 | 0.07 | 0.2 | 0.06 | 0.09
| Mean |  |  | 1.7167 | 0.1345 |  | 0.26667 | 1.5833 | 1.1333 | 2.9 | 0.085 | 0.096667 | 0.085 | 0.075 |  | 0.65 |  | 0.073333 | 0.21298 | 0.10167 | 0.21667 | 0.06 | 0.09
| Median |  |  | 1.7 | 0.1345 |  | 0.25 | 1.55 | 1.1 | 2.95 | 0.085 | 0.09 | 0.075 | 0.07 |  | 0.65 |  | 0.065 | 0.2112 | 0.105 | 0.2 | 0.06 | 0.09
| Std Dev. |  |  | 0.040825 | 0.0060133 |  | 0.12111 | 0.1472 | 0.13663 | 0.12649 | 0.0054772 | 0.012111 | 0.025884 | 0.0083666 |  | 0.054772 |  | 0.017512 | 0.030019 | 0.019408 | 0.040825 | 0 | 0.0063246
| Rel.Std.Dev. |  |  | 2.378% | 4.471% |  | 45.41% | 9.297% | 12.06% | 4.362% | 6.444% | 12.53% | 30.45% | 11.16% |  | 8.427% |  | 23.88% | 14.09% | 19.09% | 18.84% | 0% | 7.027%
| PDM3 |  |  | 1975% | 62.54% |  | 222.3% | 1813% | 1270% | 3405% | 2.719% | 16.82% | 2.719% | -9.366% |  | 685.5% |  | -11.38% | 157.4% | 22.86% | 161.8% | -27.49% | 8.761%
| Z-Score (Absolute) |  | 1.35 | NA | 1.09 | 0.1 | NA | NA | NA | NA | 0.34 | 0 | 0.34 | 0.63 | 141.7 | 15.99 | 69.45 | 0.67 | 3.36 | 0.14 | NA | 1.4 | 0.19
Results from laboratories 1.02, 1.07, 1.10, 1.11, 1.13 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 233b (Certified Value 1.6 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.84 | 1.49 | 0.80313 | 1.5 | 1.7 | 1.6 | 1.56 | 1.7 | 1.5 | 1.6 | 1.5 | 1.5 | 1.5 | 1.4 | 1.66 | 1.5 | 1.8752 | 1.6 | 1.4759 | 2 | 2.36 | 2 | 1.35 | [6 empty] | 1
| 1 | 2 | 1.99 | 1.52 | 0.75735 | 1.4 | 1.6 | 1.7 | 1.65 | 1.8 | 1.5 | 1.6 | 1.6 | 1.5 | 1.4 | 1.3 | 1.72 | 1.5 | 1.8396 | 1.6 | 1.4773 | 1.9 | 2.27 | 1.9 | 1.3 | [6 empty] | 18
| 1 | 3 | 1.88 | 1.51 | 0.72109 | 1.5 | 1.6 | 1.6 | 1.55 | 1.8 | 1.6 | 1.4 | 1.3 | 1.5 | 1.4 | 1.3 | 1.73 | 2 | 1.8276 | 1.6 | 1.4982 | 1.9 | 2.22 | 1.9 | 1.2 | [6 empty] | 16
| 1 | 4 | 1.82 | 1.56 | 0.77746 | 0.7 | 1.6 | 1.6 | 1.59 | 1.8 | 1.6 | 1.7 | 1.5 | 1.5 | 1.4 | 1.3 | 1.72 | 1.5 | 1.8465 | 1.7 | 1.4968 | 1.9 | 2.22 | 2 | 1.19 | [6 empty] | 1.5976
| 1 | 5 | 1.76 | 1.42 | 0.71551 | 1 | 1.6 | 1.7 | 1.51 | 1.7 | 1.5 | 1.7 | 1.6 | 1.6 | 1.4 | 1.4 | 1.69 | 1.5 | 1.7894 | 1.7 | 1.4394 | 2.1 | 2.13 | 2 | 1.26 | [6 empty] | 35
| 1 | 6 | 1.67 | 1.47 | 0.69236 | 0.7 | 1.5 | 1.7 | 1.58 | 1.8 | 1.5 | 1.6 | 1.5 | 1.5 | 1.5 | 1.4 | 1.69 | 1.5 | 1.8589 | 1.6 | 1.474 | 2.1 | 2.24 | 2 | 1.22
| Mean |  | 1.8267 | 1.495 | 0.74448 | 1.1333 | 1.6 | 1.65 | 1.5733 | 1.7667 | 1.5333 | 1.6 | 1.5 | 1.5167 | 1.4333 | 1.35 | 1.7017 | 1.5833 | 1.8395 | 1.6333 | 1.4769 | 1.9833 | 2.24 | 1.9667 | 1.2533
| Median |  | 1.83 | 1.5 | 0.73922 | 1.2 | 1.6 | 1.65 | 1.57 | 1.8 | 1.5 | 1.6 | 1.5 | 1.5 | 1.4 | 1.35 | 1.705 | 1.5 | 1.843 | 1.6 | 1.4766 | 1.95 | 2.23 | 2 | 1.24
| Std Dev. |  | 0.10838 | 0.047645 | 0.041944 | 0.38297 | 0.063246 | 0.054772 | 0.046762 | 0.05164 | 0.05164 | 0.10954 | 0.10954 | 0.040825 | 0.05164 | 0.054772 | 0.026394 | 0.20412 | 0.029489 | 0.05164 | 0.021286 | 0.098319 | 0.0751 | 0.05164 | 0.062503
| Rel.Std.Dev. |  | 5.933% | 3.187% | 5.634% | 33.79% | 3.953% | 3.32% | 2.972% | 2.923% | 3.368% | 6.847% | 7.303% | 2.692% | 3.603% | 4.057% | 1.551% | 12.89% | 1.603% | 3.162% | 1.441% | 4.957% | 3.353% | 2.626% | 4.987%
| PDM3 |  | 14.34% | -6.423% | -53.4% | -29.06% | 0.1496% | 3.279% | -1.52% | 10.58% | -4.023% | 0.1496% | -6.11% | -5.067% | -10.28% | -15.5% | 6.513% | -0.8936% | 15.14% | 2.236% | -7.554% | 24.14% | 40.21% | 23.1% | -21.55%
| Z-Score (Absolute) |  | 1.39 | 0.5 | 4.78 | 2.56 | 0.09 | 0.38 | 0.06 | 1.04 | 0.28 | 0.09 | 0.47 | 0.38 | 0.85 | 1.33 | 0.67 | 0 | 1.46 | 0.28 | 0.61 | 2.28 | 3.74 | 2.18 | 1.88
| Analytical results for Hg in OREAS 233b (Indicative Value < 1 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab |  |  |  | 1
| No | No | 1.01: 1.05; 1.02: 1.25; 1.04: 1.30 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*OES/MS; 1.04: 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 1.01: < 1; 1.02: < 0.1; 1.04: 0.0401 |  |  |  | 1
| 1 | 2 | 1.01: < 1; 1.02: < 0.1; 1.04: 0.0303 |  |  |  | 5
| 1 | 3 | 1.01: < 1; 1.02: < 0.1; 1.04: 0.0208 |  |  |  | 16
| 1 | 4 | 1.01: < 1; 1.02: < 0.1; 1.04: < 0.01 |  |  |  | < 1
| 1 | 5 | 1.01: < 1; 1.02: 0.393; 1.04: < 0.01 |  |  |  | 9
| 1 | 6 | 1.01: < 1; 1.02: < 0.1; 1.04: < 0.01
| Mean |  | 1.02: 0.393; 1.04: 0.0304
| Median |  | 1.02: 0.393; 1.04: 0.0303
| Std Dev. |  | 1.04: 0.0096504
| Rel.Std.Dev. |  | 1.04: 31.74%
PDM3
| Z-Score (Absolute) |  | 1.01: 2.96; 1.02: 0; 1.04: 0.67
| Analytical results for Ho in OREAS 233b (Certified Value 0.74 ppm) | [30 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab; 1.11: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.08; 1.08: 1.24; 1.10: 1.27; 1.11: 1.28 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.07: 4A*OES; 1.08: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.01: 0.77; 1.02: 0.63; 1.04: 0.69071; 1.05: 0.8; 1.07: 0.75; 1.08: 0.78; 1.10: 0.75; 1.11: 0.80199 |  |  |  | 1
| 1 | 2 | 1.01: 0.77; 1.02: 0.72; 1.04: 0.71433; 1.05: 0.7; 1.07: 0.71; 1.08: 0.78; 1.10: 0.8; 1.11: 0.77916 |  |  |  | 19
| 1 | 3 | 1.01: 0.77; 1.02: 0.68; 1.04: 0.68138; 1.05: 0.7; 1.07: 0.72; 1.08: 0.8; 1.10: 0.75; 1.11: 0.78161 |  |  |  | 16
| 1 | 4 | 1.01: 0.79; 1.02: 0.76; 1.04: 0.72322; 1.05: 0.7; 1.07: 0.72; 1.08: 0.78; 1.10: 0.75; 1.11: 0.7919 |  |  |  | 0.74299
| 1 | 5 | 1.01: 0.78; 1.02: 0.69; 1.04: 0.73502; 1.05: 0.7; 1.07: 0.72; 1.08: 0.76; 1.10: 0.7; 1.11: 0.75492 |  |  |  | 36
| 1 | 6 | 1.01: 0.77; 1.02: 0.65; 1.04: 0.68911; 1.05: 0.7; 1.07: 0.68; 1.08: 0.77; 1.10: 0.75; 1.11: 0.81236
| Mean |  | 1.01: 0.775; 1.02: 0.68833; 1.04: 0.70563; 1.05: 0.71667; 1.07: 0.71667; 1.08: 0.77833; 1.10: 0.75; 1.11: 0.78699
| Median |  | 1.01: 0.77; 1.02: 0.685; 1.04: 0.70252; 1.05: 0.7; 1.07: 0.72; 1.08: 0.78; 1.10: 0.75; 1.11: 0.78675
| Std Dev. |  | 1.01: 0.0083666; 1.02: 0.047081; 1.04: 0.021595; 1.05: 0.040825; 1.07: 0.022509; 1.08: 0.013292; 1.10: 0.031623; 1.11: 0.020057
| Rel.Std.Dev. |  | 1.01: 1.08%; 1.02: 6.84%; 1.04: 3.06%; 1.05: 5.696%; 1.07: 3.141%; 1.08: 1.708%; 1.10: 4.216%; 1.11: 2.549%
| PDM3 |  | 1.01: 4.308%; 1.02: -7.357%; 1.04: -5.029%; 1.05: -3.543%; 1.07: -3.543%; 1.08: 4.756%; 1.10: 0.9431%; 1.11: 5.922%
| Z-Score (Absolute) |  | 1.01: 0.51; 1.02: 1.25; 1.04: 0.9; 1.05: NA; 1.07: 0.67; 1.08: 0.57; 1.10: 0; 1.11: 0.75
| Analytical results for In in OREAS 233b (Certified Value 0.073 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.08 | 0.08 | 0.054157 | < 0.1 | 0.06 | 0.08 | 0.08 | 0.082 | 0.081 | 0.067 | 0.078 | < 10 | 0.07 | 0.07 | 0.07 | 31.95 | < 0.1 | 0.06717 | 0.039 | 0.0924 | 0.078 | 0.08 | 0.075 | 0.066 | [5 empty] | 1
| 1 | 2 | 0.08 | 0.07 | 0.060367 | < 0.1 | 0.07 | 0.08 | 0.08 | 0.079 | 0.074 | 0.069 | 0.081 | < 10 | 0.07 | 0.07 | 0.069 | 32.76 | < 0.1 | 0.06473 | 0.049 | 0.0865 | 0.077 | 0.07 | 0.075 | 0.062 | [5 empty] | 20
| 1 | 3 | 0.09 | 0.07 | 0.056889 | < 0.1 | 0.06 | 0.07 | 0.08 | 0.082 | 0.067 | 0.069 | 0.081 | < 10 | 0.07 | 0.06 | 0.076 | 27.65 | < 0.1 | 0.06246 | 0.051 | 0.0854 | 0.072 | 0.07 | 0.077 | 0.061 | [5 empty] | 16
| 1 | 4 | 0.08 | 0.08 | 0.059745 | < 0.1 | 0.08 | 0.08 | 0.09 | 0.078 | 0.079 | 0.073 | 0.075 | < 10 | 0.06 | 0.07 | 0.07 | 29.11 | < 0.1 | 0.06636 | 0.056 | 0.0823 | 0.069 | 0.07 | 0.083 | 0.057 | [5 empty] | 0.073004
| 1 | 5 | 0.08 | 0.07 | 0.056889 | < 0.1 | 0.07 | 0.07 | 0.09 | 0.074 | 0.08 | 0.069 | 0.086 | < 10 | 0.07 | 0.07 | 0.072 | 27.32 | < 0.1 | 0.06537 | 0.051 | 0.0775 | 0.08 | 0.07 | 0.082 | 0.063 | [5 empty] | 37
| 1 | 6 | 0.08 | 0.08 | 0.052293 | < 0.1 | 0.06 | 0.07 | 0.08 | 0.084 | 0.072 | 0.069 | 0.086 | < 10 | 0.07 | 0.07 | 0.071 | 25.8 | < 0.1 | 0.06966 | 0.045 | 0.0813 | 0.072 | 0.07 | 0.081 | 0.058
| Mean |  | 0.081667 | 0.075 | 0.056723 |  | 0.066667 | 0.075 | 0.083333 | 0.079833 | 0.0755 | 0.069333 | 0.081167 |  | 0.068333 | 0.068333 | 0.071333 | 29.0983 |  | 0.065958 | 0.0485 | 0.084233 | 0.074667 | 0.071667 | 0.078833 | 0.061167
| Median |  | 0.08 | 0.075 | 0.056889 |  | 0.065 | 0.075 | 0.08 | 0.0805 | 0.0765 | 0.069 | 0.081 |  | 0.07 | 0.07 | 0.0705 | 28.38 |  | 0.065865 | 0.05 | 0.08385 | 0.0745 | 0.07 | 0.079 | 0.0615
| Std Dev. |  | 0.0040825 | 0.0054772 | 0.0031204 |  | 0.008165 | 0.0054772 | 0.005164 | 0.0036009 | 0.0054681 | 0.0019664 | 0.0043551 |  | 0.0040825 | 0.0040825 | 0.0025033 | 2.7451 |  | 0.0024259 | 0.0058566 | 0.0051161 | 0.004274 | 0.0040825 | 0.0036009 | 0.0033116
| Rel.Std.Dev. |  | 4.999% | 7.303% | 5.501% |  | 12.25% | 7.303% | 6.197% | 4.511% | 7.243% | 2.836% | 5.366% |  | 5.974% | 5.974% | 3.509% | 9.434% |  | 3.678% | 12.08% | 6.074% | 5.724% | 5.696% | 4.568% | 5.414%
| PDM3 |  | 11.87% | 2.734% | -22.3% |  | -8.681% | 2.734% | 14.15% | 9.354% | 3.419% | -5.028% | 11.18% |  | -6.398% | -6.398% | -2.289% | 39758% |  | -9.651% | -33.57% | 15.38% | 2.277% | -1.832% | 7.985% | -16.21%
| Z-Score (Absolute) |  | 0.83 | 0.18 | 1.6 | 2.25 | 0.63 | 0.18 | 0.99 | 0.65 | 0.23 | 0.37 | 0.78 | 478.88 | 0.47 | 0.47 | 0.18 | 2821.18 | 2.25 | 0.7 | 2.4 | 1.08 | 0.15 | 0.15 | 0.55 | 1.17
| Analytical results for K in OREAS 233b (Certified Value 0.541 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.57 | 0.5366 | 0.52422 | 0.52333 | 0.56 | 0.5836 | 0.529 | 0.504 | 0.54 | 0.56 | 0.55 | 0.53 | 0.51 | 0.51 | 0.54 | 0.537 | 0.5 | 0.57 | 0.586 | 0.52173 | 0.52 | 0.54 | 0.5713 | 0.53 | 0.61 | 0.57 | 0.54 |  |  | 1
| 1 | 2 | 0.56 | 0.5313 | 0.52171 | 0.5306 | 0.54 | 0.5861 | 0.516 | 0.5151 | 0.54 | 0.58 | 0.56 | 0.52 | 0.53 | 0.51 | 0.55 | 0.536 | 0.5 | 0.58 | 0.594 | 0.52705 | 0.515 | 0.53 | 0.5869 | 0.54 | 0.59 | 0.58 | 0.54 |  |  | 0
| 1 | 3 | 0.57 | 0.5382 | 0.52128 | 0.52379 | 0.53 | 0.5819 | 0.524 | 0.5022 | 0.54 | 0.56 | 0.56 | 0.52 | 0.52 | 0.52 | 0.54 | 0.532 | 0.49 | 0.56 | 0.701 | 0.54382 | 0.515 | 0.53 | 0.5856 | 0.53 | 0.6 | 0.57 | 0.53 |  |  | 16
| 1 | 4 | 0.56 | 0.5325 | 0.52759 | 0.53282 | 0.52 | 0.5819 | 0.528 | 0.5069 | 0.53 | 0.59 | 0.56 | 0.52 | 0.51 | 0.52 | 0.55 | 0.542 | 0.5 | 0.56 | 0.677 | 0.53695 | 0.515 | 0.53 | 0.585 | 0.54 | 0.59 | 0.57 | 0.52 |  |  | 0.54135
| 1 | 5 | 0.55 | 0.5259 | 0.52593 | 0.5397 | 0.52 | 0.577 | 0.524 | 0.4949 | 0.53 | 0.58 | 0.57 | 0.53 | 0.5 | 0.52 | 0.57 | 0.536 | 0.51 | 0.6 | 0.788 | 0.52806 | 0.515 | 0.53 | 0.5686 | 0.54 | 0.57 | 0.57 | 0.54 |  |  | 38
| 1 | 6 | 0.57 | 0.5342 | 0.52321 | 0.50327 | 0.51 | 0.5869 | 0.54 | 0.5121 | 0.52 | 0.58 | 0.56 | 0.53 | 0.54 | 0.52 | 0.56 | 0.543 | 0.5 | 0.56 | 0.732 | 0.54565 | 0.515 | 0.53 | 0.5797 | 0.54 | 0.59 | 0.57 | 0.53
| Mean |  | 0.56333 | 0.53312 | 0.52399 | 0.52559 | 0.53 | 0.5829 | 0.52683 | 0.50587 | 0.53333 | 0.575 | 0.56 | 0.525 | 0.51833 | 0.51667 | 0.55167 | 0.53767 | 0.5 | 0.57167 | 0.67967 | 0.53388 | 0.51583 | 0.53167 | 0.57952 | 0.53667 | 0.59167 | 0.57167 | 0.53333
| Median |  | 0.565 | 0.53335 | 0.52372 | 0.52719 | 0.525 | 0.58275 | 0.526 | 0.50545 | 0.535 | 0.58 | 0.56 | 0.525 | 0.515 | 0.52 | 0.55 | 0.5365 | 0.5 | 0.565 | 0.689 | 0.5325 | 0.515 | 0.53 | 0.58235 | 0.54 | 0.59 | 0.57 | 0.535
| Std Dev. |  | 0.008165 | 0.0043568 | 0.0024477 | 0.012514 | 0.017889 | 0.0035648 | 0.0079099 | 0.0072439 | 0.008165 | 0.012247 | 0.0063246 | 0.0054772 | 0.01472 | 0.005164 | 0.01169 | 0.0041312 | 0.0063246 | 0.016021 | 0.078803 | 0.0097472 | 0.0020412 | 0.0040825 | 0.0078518 | 0.005164 | 0.013292 | 0.0040825 | 0.008165
| Rel.Std.Dev. |  | 1.449% | 0.8172% | 0.4671% | 2.381% | 3.375% | 0.6116% | 1.501% | 1.432% | 1.531% | 2.13% | 1.129% | 1.043% | 2.84% | 0.9995% | 2.119% | 0.7684% | 1.265% | 2.802% | 11.59% | 1.826% | 0.3957% | 0.7679% | 1.355% | 0.9622% | 2.246% | 0.7141% | 1.531%
| PDM3 |  | 4.06% | -1.522% | -3.207% | -2.912% | -2.097% | 7.675% | -2.682% | -6.555% | -1.482% | 6.215% | 3.444% | -3.021% | -4.252% | -4.56% | 1.905% | -0.6811% | -7.639% | 5.599% | 25.55% | -1.381% | -4.714% | -1.789% | 7.05% | -0.8658% | 9.294% | 5.599% | -1.482%
| Z-Score (Absolute) |  | 1.21 | 0.01 | 0.38 | 0.31 | 0.13 | 2.01 | 0.26 | 1.11 | 0 | 1.69 | 1.08 | 0.34 | 0.61 | 0.67 | 0.74 | 0.18 | 1.35 | 1.55 | 5.92 | 0.02 | 0.71 | 0.07 | 1.87 | 0.13 | 2.36 | 1.55 | 0
| Analytical results for La in OREAS 233b (Certified Value 5.3 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 5.8 | 5.39 | 4.57 | 5.43 | 5.1394 | 5.4 | 6 | 5.22 | 5 | 8 | 5.1 | 5.5 | 5.1 | 5 | 5 | 4.9 | 5.2 | 6 | 5.2577 | 5.5 | 5.4658 | 6.4 | 5.6 | 6 | 5.2 |  |  |  |  | 1
| 1 | 2 | 6.4 | 5.28 | 4.99 | 5.49 | 5.1825 | 5.3 | 6 | 5.14 | 5 | 9 | 5.4 | 5.3 | 5.5 | 5.3 | 6 | 4.8 | 5.2 | 6.5 | 5.2805 | 5.5 | 5.4963 | 6.3 | 5.6 | 6 | 4.8 |  |  |  |  | 21
| 1 | 3 | 5.8 | 5.26 | 4.89 | 5.45 | 5.1697 | 5.2 | 6 | 5.15 | 5 | 8 | 5.3 | 5 | 5.1 | 5.5 | 5 | 4.8 | 5.1 | 6 | 5.2535 | 5.4 | 5.5019 | 6.3 | 5.5 | 6 | 4.7 |  |  |  |  | 16
| 1 | 4 | 5.6 | 5.35 | 5.21 | 5.45 | 5.2228 | 5.1 | 6 | 5.36 | 5 | 8 | 5.3 | 5.5 | 5.4 | 5.4 | 6 | 4.9 | 5 | 6.5 | 5.2455 | 5.4 | 5.5444 | 6.4 | 5.5 | 6.3 | 4.6 |  |  |  |  | 5.301
| 1 | 5 | 5.6 | 5.23 | 4.97 | 5.42 | 5.3557 | 5.1 | 6 | 5.09 | 5 | 9 | 5.3 | 5.4 | 5.5 | 5.5 | 6 | 4.9 | 5.4 | 6 | 5.042 | 5.5 | 5.3943 | 6.5 | 5.3 | 6.2 | 4.8 |  |  |  |  | 39
| 1 | 6 | 5.8 | 5.35 | 4.73 | 5.41 | 5.0788 | 5.2 | 6 | 5.2 | 5 | 8 | 5.3 | 5.7 | 5.2 | 5.4 | 6 | 4.8 | 5 | 6 | 5.3006 | 5.4 | 5.3903 | 6.4 | 5.5 | 6.3 | 4.6
| Mean |  | 5.8333 | 5.31 | 4.8933 | 5.4417 | 5.1915 | 5.2167 | 6 | 5.1933 | 5 | 8.3333 | 5.2833 | 5.4 | 5.3 | 5.35 | 5.6667 | 4.85 | 5.15 | 6.1667 | 5.2299 | 5.45 | 5.4655 | 6.3833 | 5.5 | 6.1333 | 4.7833
| Median |  | 5.8 | 5.315 | 4.93 | 5.44 | 5.1761 | 5.2 | 6 | 5.175 | 5 | 8 | 5.3 | 5.45 | 5.3 | 5.4 | 6 | 4.85 | 5.15 | 6 | 5.2556 | 5.45 | 5.481 | 6.4 | 5.5 | 6.1 | 4.75
| Std Dev. |  | 0.29439 | 0.06229 | 0.22214 | 0.028577 | 0.093699 | 0.1169 | 0 | 0.093737 | 0 | 0.5164 | 0.098319 | 0.23664 | 0.18974 | 0.18708 | 0.5164 | 0.054772 | 0.15166 | 0.2582 | 0.09428 | 0.054772 | 0.062005 | 0.075277 | 0.10954 | 0.15055 | 0.22286
| Rel.Std.Dev. |  | 5.047% | 1.173% | 4.54% | 0.5252% | 1.805% | 2.241% | 0% | 1.805% | 0% | 6.197% | 1.861% | 4.382% | 3.58% | 3.497% | 9.113% | 1.129% | 2.945% | 4.187% | 1.803% | 1.005% | 1.134% | 1.179% | 1.992% | 2.455% | 4.659%
| PDM3 |  | 10.04% | 0.1692% | -7.691% | 2.653% | -2.067% | -1.591% | 13.19% | -2.032% | -5.679% | 57.2% | -0.3339% | 1.867% | -0.01945% | 0.9238% | 6.897% | -8.508% | -2.849% | 16.33% | -1.341% | 2.81% | 3.103% | 20.42% | 3.753% | 15.7% | -9.766%
| Z-Score (Absolute) |  | 2.52 | 0 | 2.01 | 0.63 | 0.57 | 0.45 | NA | 0.56 | NA | NA | 0.13 | 0.43 | 0.05 | 0.19 | NA | 2.22 | 0.77 | 4.13 | 0.39 | 0.67 | 0.75 | 5.17 | 0.92 | 3.97 | 2.54
Results from laboratories 1.08, 1.12, 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 233b (Certified Value 12.2 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | < 1 | 12.4 | 10.31 | 8.1411 | 12.933 | 12 | 6 | 13.4 | 12.3 | 12 | 14 | 13.6 | 12.2 | 12.6 | 12.7 | < 5 | 11.7 | 10.4 | 11.2 | 13 | 9.8533 | 12.9 | 10.2414 | 12.3 | 16 | 12.6 | 11.7 |  |  | 1
| 1 | 2 | < 1 | 12.3 | 10.02 | 8.1599 | 13.2353 | 12 | 6 | 13.1 | 11.8 | 13 | 13.2 | 14.4 | 12 | 13.3 | 12.5 | < 5 | 11.2 | 10.5 | 11.5 | 13 | 9.5896 | 12.8 | 10.491 | 12.5 | 15 | 12.4 | 11.9 |  |  | 0
| 1 | 3 | < 1 | 12 | 10.57 | 8.2488 | 12.956 | 12 | 6 | 13.6 | 12.2 | 11 | 13.2 | 14.2 | 11.9 | 12.7 | 12.7 | < 5 | 11.9 | 10.4 | 11.9 | 13 | 9.3989 | 12.8 | 10.4442 | 12.3 | 15 | 12.7 | 10.8 |  |  | 16
| 1 | 4 | < 1 | 12.2 | 15.1 | 8.6064 | 13.0972 | 12 | 6 | 13.3 | 12.2 | 12 | 13.3 | 14.2 | 11.8 | 12.6 | 12.7 | < 5 | 11.1 | 10.7 | 11.6 | 13 | 9.8779 | 12.6 | 10.5144 | 12.5 | 15 | 13.1 | 10.7 |  |  | 12.192
| 1 | 5 | < 1 | 11.6 | 14.09 | 8.251 | 13.3431 | 12 | 6 | 13.3 | 11.8 | 12 | 15.3 | 14.3 | 12 | 12.4 | 12.5 | < 5 | 11.7 | 10.7 | 11.5 | 13 | 9.6627 | 12.5 | 10.2414 | 12.5 | 15 | 13.1 | 11.3 |  |  | 40
| 1 | 6 | < 1 | 12.1 | 13.92 | 7.9871 | 12.1234 | 11 | 6 | 13.8 | 11.9 | 12 | 14.9 | 14.2 | 12.4 | 13.6 | 12.8 | < 5 | 11.8 | 10.5 | 11.2 | 13 | 10.2092 | 12.6 | 10.3116 | 12.5 | 15 | 12.7 | 10.9
| Mean |  |  | 12.1 | 12.335 | 8.2324 | 12.948 | 11.8333 | 6 | 13.4167 | 12.0333 | 12 | 13.9833 | 14.15 | 12.05 | 12.8667 | 12.65 |  | 11.5667 | 10.5333 | 11.4833 | 13 | 9.7653 | 12.7 | 10.374 | 12.4333 | 15.1667 | 12.7667 | 11.2167
| Median |  |  | 12.15 | 12.245 | 8.2043 | 13.0266 | 12 | 6 | 13.35 | 12.05 | 12 | 13.65 | 14.2 | 12 | 12.65 | 12.7 |  | 11.7 | 10.5 | 11.5 | 13 | 9.758 | 12.7 | 10.3779 | 12.5 | 15 | 12.7 | 11.1
| Std Dev. |  |  | 0.28284 | 2.2721 | 0.207 | 0.43386 | 0.40825 | 0 | 0.24833 | 0.22509 | 0.63246 | 0.92394 | 0.28107 | 0.21679 | 0.47188 | 0.12247 |  | 0.33267 | 0.13663 | 0.26394 | 0 | 0.28039 | 0.15492 | 0.12441 | 0.10328 | 0.40825 | 0.28048 | 0.49967
| Rel.Std.Dev. |  |  | 2.338% | 18.42% | 2.514% | 3.351% | 3.45% | 0% | 1.851% | 1.871% | 5.27% | 6.607% | 1.986% | 1.799% | 3.667% | 0.9682% |  | 2.876% | 1.297% | 2.299% | 0% | 2.871% | 1.22% | 1.199% | 0.8307% | 2.692% | 2.197% | 4.455%
| PDM3 |  |  | -0.7542% | 1.173% | -32.48% | 6.201% | -2.941% | -50.79% | 10.05% | -1.301% | -1.574% | 14.69% | 16.06% | -1.164% | 5.534% | 3.757% |  | -5.129% | -13.6% | -5.812% | 6.628% | -19.9% | 4.167% | -14.91% | 1.98% | 24.4% | 4.714% | -7.999%
| Z-Score (Absolute) |  | 9.46 | 0.02 | 0.21 | 3.14 | 0.71 | NA | NA | 1.1 | 0.03 | NA | 1.56 | 1.7 | 0.02 | 0.65 | 0.47 | 7.83 | 0.42 | 1.26 | 0.48 | NA | 1.89 | 0.51 | 1.39 | 0.29 | NA | 0.57 | 0.7
Results from laboratories 1.07, 1.08, 1.12, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 233b (Certified Value 0.32 ppm) | [30 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28; 1.11: 1.32 |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.01: 0.35; 1.02: 0.32; 1.03: 0.25; 1.04: 4.2251; 1.05: 0.3; 1.07: 0.33; 1.08: 0.5; 1.10: 0.33908; 1.11: 0.3 |  |  | 1
| 1 | 2 | 1.01: 0.35; 1.02: 0.31; 1.03: 0.28; 1.04: 4.2084; 1.05: 0.3; 1.07: 0.33; 1.08: 0.5; 1.10: 0.32495; 1.11: 0.3 |  |  | 22
| 1 | 3 | 1.01: 0.34; 1.02: 0.32; 1.03: 0.27; 1.04: 4.2302; 1.05: 0.3; 1.07: 0.34; 1.08: 0.5; 1.10: 0.32266; 1.11: 0.29 |  |  | 16
| 1 | 4 | 1.01: 0.35; 1.02: 0.32; 1.03: 0.28; 1.04: 4.1672; 1.05: 0.3; 1.07: 0.33; 1.08: 0.5; 1.10: 0.33769; 1.11: 0.29 |  |  | 0.32391
| 1 | 5 | 1.01: 0.35; 1.02: 0.3; 1.03: 0.28; 1.04: 4.1838; 1.05: 0.3; 1.07: 0.33; 1.08: 0.5; 1.10: 0.31839; 1.11: 0.28 |  |  | 41
| 1 | 6 | 1.01: 0.34; 1.02: 0.32; 1.03: 0.26; 1.04: 4.2189; 1.05: 0.3; 1.07: 0.34; 1.08: 0.5; 1.10: 0.34463; 1.11: 0.3
| Mean |  | 1.01: 0.34667; 1.02: 0.315; 1.03: 0.27; 1.04: 4.2056; 1.05: 0.3; 1.07: 0.33333; 1.08: 0.5; 1.10: 0.33123; 1.11: 0.29333
| Median |  | 1.01: 0.35; 1.02: 0.32; 1.03: 0.275; 1.04: 4.2137; 1.05: 0.3; 1.07: 0.33; 1.08: 0.5; 1.10: 0.33132; 1.11: 0.295
| Std Dev. |  | 1.01: 0.005164; 1.02: 0.0083666; 1.03: 0.012649; 1.04: 0.024957; 1.05: 0; 1.07: 0.005164; 1.08: 0; 1.10: 0.010589; 1.11: 0.008165
| Rel.Std.Dev. |  | 1.01: 1.49%; 1.02: 2.656%; 1.03: 4.685%; 1.04: 0.5934%; 1.05: 0%; 1.07: 1.549%; 1.08: 0%; 1.10: 3.197%; 1.11: 2.784%
| PDM3 |  | 1.01: 7.025%; 1.02: -2.752%; 1.03: -16.64%; 1.04: 1198%; 1.05: -7.383%; 1.07: 2.908%; 1.08: 54.36%; 1.10: 2.26%; 1.11: -9.441%
| Z-Score (Absolute) |  | 1.01: 0.64; 1.02: 0.67; 1.03: 2.54; 1.04: 160.94; 1.05: NA; 1.07: 0.09; 1.08: NA; 1.10: 0; 1.11: 1.57
Results from laboratories 1.07 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 233b (Certified Value 3.87 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.8 | 3.7684 | 3.9511 | 3.7123 | 3.86 | 4.1973 | 3.94 | 3.7281 | 3.67 | 3.38 | 4.07 | 3.94 | 3.77 | 3.85 | 3.79 | 3.93 | 3.86 | 3.774 | 0.825 | 3.1224 | 4.019 | 3.94 | 3.7848 | 3.88 | 3.09 | 3.88 | 4 |  |  | 1
| 1 | 2 | 3.79 | 3.7298 | 3.9551 | 3.7696 | 3.7 | 4.2094 | 3.84 | 3.8118 | 3.69 | 3.45 | 4.16 | 3.81 | 3.86 | 3.83 | 3.92 | 3.943 | 3.91 | 3.794 | 1.11 | 3.1359 | 4.024 | 3.93 | 3.8696 | 3.95 | 3.06 | 3.93 | 3.9 |  |  | 0
| 1 | 3 | 3.77 | 3.7429 | 3.9851 | 3.7098 | 3.68 | 4.1672 | 3.91 | 3.7981 | 3.67 | 3.31 | 4.18 | 3.8 | 3.72 | 3.88 | 3.83 | 3.889 | 3.89 | 3.693 | 0.838 | 3.2784 | 4.028 | 3.92 | 3.8678 | 3.84 | 3.07 | 3.88 | 4 |  |  | 16
| 1 | 4 | 3.73 | 3.6772 | 3.9568 | 3.7639 | 3.64 | 4.1853 | 3.91 | 3.7729 | 3.66 | 3.44 | 4.17 | 3.8 | 3.75 | 3.86 | 3.87 | 3.992 | 3.88 | 3.676 | 0.953 | 3.1639 | 4.021 | 3.88 | 3.8827 | 3.95 | 3.06 | 3.9 | 4 |  |  | 3.8702
| 1 | 5 | 3.68 | 3.6828 | 3.9553 | 3.8318 | 3.52 | 4.1491 | 3.9 | 3.7398 | 3.64 | 3.41 | 4.2 | 3.89 | 3.72 | 3.81 | 4.03 | 3.915 | 3.87 | 3.961 | 0.389 | 3.1128 | 4.0005 | 3.87 | 3.7979 | 3.91 | 2.95 | 3.9 | 4 |  |  | 42
| 1 | 6 | 3.78 | 3.7207 | 3.9371 | 3.5552 | 3.56 | 4.1913 | 4.01 | 3.7998 | 3.63 | 3.42 | 4.14 | 3.77 | 4.01 | 3.88 | 3.87 | 3.979 | 3.88 | 3.714 | 0.446 | 3.224 | 4.0375 | 3.93 | 3.8209 | 3.94 | 3.03 | 3.89 | 4.1
| Mean |  | 3.7583 | 3.7203 | 3.9567 | 3.7238 | 3.66 | 4.1833 | 3.9183 | 3.7751 | 3.66 | 3.4017 | 4.1533 | 3.835 | 3.805 | 3.8517 | 3.885 | 3.9413 | 3.8817 | 3.7687 | 0.76017 | 3.1729 | 4.0217 | 3.9117 | 3.8373 | 3.9117 | 3.0433 | 3.8967 | 4
| Median |  | 3.775 | 3.7252 | 3.9552 | 3.7381 | 3.66 | 4.1883 | 3.91 | 3.7855 | 3.665 | 3.415 | 4.165 | 3.805 | 3.76 | 3.855 | 3.87 | 3.9365 | 3.88 | 3.744 | 0.8315 | 3.1499 | 4.0225 | 3.925 | 3.8444 | 3.925 | 3.06 | 3.895 | 4
| Std Dev. |  | 0.04535 | 0.035153 | 0.015636 | 0.093945 | 0.12 | 0.02179 | 0.055648 | 0.034485 | 0.021909 | 0.051153 | 0.045461 | 0.065345 | 0.11292 | 0.027869 | 0.083367 | 0.038857 | 0.017224 | 0.10485 | 0.28506 | 0.065359 | 0.012262 | 0.029269 | 0.041503 | 0.04446 | 0.049666 | 0.018619 | 0.063246
| Rel.Std.Dev. |  | 1.207% | 0.9449% | 0.3952% | 2.523% | 3.279% | 0.5209% | 1.42% | 0.9135% | 0.5986% | 1.504% | 1.095% | 1.704% | 2.968% | 0.7236% | 2.146% | 0.9859% | 0.4437% | 2.782% | 37.5% | 2.06% | 0.3049% | 0.7482% | 1.082% | 1.137% | 1.632% | 0.4778% | 1.581%
| PDM3 |  | -2.891% | -3.874% | 2.235% | -3.784% | -5.432% | 8.088% | 1.243% | -2.459% | -5.432% | -12.11% | 7.315% | -0.9105% | -1.686% | -0.4798% | 0.3814% | 1.837% | 0.2953% | -2.624% | -80.36% | -18.02% | 3.913% | 1.07% | -0.8509% | 1.07% | -21.37% | 0.6829% | 3.353%
| Z-Score (Absolute) |  | 0.51 | 0.76 | 0.77 | 0.74 | 1.15 | 2.24 | 0.53 | 0.4 | 1.15 | 2.82 | 2.05 | 0.01 | 0.21 | 0.09 | 0.31 | 0.67 | 0.29 | 0.44 | 19.95 | 4.31 | 1.2 | 0.48 | 0 | 0.48 | 5.15 | 0.38 | 1.05
| Analytical results for Mn in OREAS 233b (Certified Value 0.147 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.15 | 0.1504 | 0.15399 | 0.14282 | 0.154 | 0.1657 | 0.148 | 0.1471 | 0.1354 | 0.1324 | 0.148 | 0.1505 | 0.1475 | 0.1475 | 0.1457 | 0.1526 | 0.1484 | 0.1469 | 0.087 | 0.13355 | 0.146 | 0.1435 | 0.12762 | 0.15 | 0.152 | 0.1438 | 0.144 |  |  | 1
| 1 | 2 | 0.15 | 0.1534 | 0.15534 | 0.14464 | 0.152 | 0.1665 | 0.144 | 0.1479 | 0.1353 | 0.1632 | 0.152 | 0.146 | 0.15 | 0.145 | 0.149 | 0.1502 | 0.1503 | 0.1477 | 0.089 | 0.13465 | 0.145 | 0.1455 | 0.1301 | 0.153 | 0.15 | 0.1464 | 0.143 |  |  | 23
| 1 | 3 | 0.15 | 0.1524 | 0.15501 | 0.14277 | 0.149 | 0.1657 | 0.147 | 0.148 | 0.1354 | 0.1655 | 0.1525 | 0.1455 | 0.148 | 0.1485 | 0.1473 | 0.1502 | 0.1496 | 0.1448 | 0.084 | 0.14052 | 0.145 | 0.1435 | 0.13021 | 0.15 | 0.1485 | 0.1445 | 0.139 |  |  | 16
| 1 | 4 | 0.15 | 0.1492 | 0.15515 | 0.145 | 0.142 | 0.1657 | 0.147 | 0.1471 | 0.1358 | 0.1313 | 0.154 | 0.145 | 0.1475 | 0.1475 | 0.1477 | 0.1523 | 0.1494 | 0.1429 | 0.085 | 0.13796 | 0.146 | 0.142 | 0.13021 | 0.153 | 0.1476 | 0.1453 | 0.14 |  |  | 0.14715
| 1 | 5 | 0.15 | 0.1469 | 0.15629 | 0.14726 | 0.14 | 0.165 | 0.147 | 0.1452 | 0.1354 | 0.1328 | 0.1515 | 0.1475 | 0.146 | 0.1455 | 0.1516 | 0.1521 | 0.1496 | 0.1541 | 0.078 | 0.13588 | 0.146 | 0.1435 | 0.12848 | 0.151 | 0.1456 | 0.145 | 0.143 |  |  | 43
| 1 | 6 | 0.15 | 0.1522 | 0.15342 | 0.13713 | 0.141 | 0.1665 | 0.15 | 0.1481 | 0.1356 | 0.1629 | 0.1505 | 0.146 | 0.1565 | 0.1465 | 0.146 | 0.1511 | 0.1489 | 0.1454 | 0.081 | 0.14072 | 0.1455 | 0.1445 | 0.12838 | 0.152 | 0.1498 | 0.1442 | 0.145
| Mean |  | 0.15 | 0.15075 | 0.15487 | 0.14327 | 0.14633 | 0.16585 | 0.14717 | 0.14723 | 0.13548 | 0.14802 | 0.15142 | 0.14675 | 0.14925 | 0.14675 | 0.14788 | 0.15142 | 0.14937 | 0.14697 | 0.084 | 0.13721 | 0.14558 | 0.14375 | 0.12917 | 0.1515 | 0.14892 | 0.14487 | 0.14233
| Median |  | 0.15 | 0.1513 | 0.15508 | 0.14373 | 0.1455 | 0.1657 | 0.147 | 0.1475 | 0.1354 | 0.14785 | 0.15175 | 0.146 | 0.14775 | 0.147 | 0.1475 | 0.1516 | 0.1495 | 0.14615 | 0.0845 | 0.13692 | 0.14575 | 0.1435 | 0.12929 | 0.1515 | 0.14915 | 0.14475 | 0.143
| Std Dev. |  | 0 | 0.0024164 | 0.0010219 | 0.0034334 | 0.0060882 | 0.00057184 | 0.0019408 | 0.0010912 | 0.00018348 | 0.017393 | 0.0020351 | 0.0020187 | 0.0037782 | 0.0013323 | 0.0021794 | 0.0010685 | 0.0006532 | 0.0038759 | 0.004 | 0.0030165 | 0.0004916 | 0.0011726 | 0.001144 | 0.0013784 | 0.0022076 | 0.00092448 | 0.0023381
| Rel.Std.Dev. |  | 0% | 1.603% | 0.6599% | 2.396% | 4.161% | 0.3448% | 1.319% | 0.7411% | 0.1354% | 11.75% | 1.344% | 1.376% | 2.531% | 0.9079% | 1.474% | 0.7057% | 0.4373% | 2.637% | 4.762% | 2.198% | 0.3377% | 0.8157% | 0.8857% | 0.9098% | 1.482% | 0.6382% | 1.643%
| PDM3 |  | 1.936% | 2.446% | 5.243% | -2.638% | -0.5557% | 12.71% | 0.01061% | 0.05591% | -7.929% | 0.5882% | 2.899% | -0.2725% | 1.426% | -0.2725% | 0.4976% | 2.899% | 1.506% | -0.1253% | -42.92% | -6.754% | -1.065% | -2.311% | -12.22% | 2.955% | 1.2% | -1.552% | -3.274%
| Z-Score (Absolute) |  | 0.67 | 0.85 | 1.83 | 0.93 | 0.2 | 4.45 | 0 | 0.02 | 2.78 | 0.2 | 1.01 | 0.1 | 0.5 | 0.1 | 0.17 | 1.01 | 0.52 | 0.05 | 15.03 | 2.37 | 0.38 | 0.81 | 4.28 | 1.03 | 0.42 | 0.55 | 1.15
| Analytical results for Mo in OREAS 233b (Certified Value 2.34 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.66 | 2.4 | 2.3 | 2.6166 | 2.1544 | 2 | 2.89 | 2.2 | 2.3 | 2 | 2 | 2.33 | 2.45 | 2.21 | 2.12 | 2 | 1.9 | 2 | 2.25 | 3.01 | 2 | 2.3184 | 2.38 | 2.2934 | 2.49 | 2.8 | 2.39 | 2.1 |  | 1
| 1 | 2 | 2.63 | 2.3 | 2.5 | 2.3756 | 2.2473 | 1 | 2.84 | 2.3 | 2.4 | 2 | 1.8 | 2.47 | 2.39 | 2.44 | 2.37 | < 2 | 1.9 | 2.1 | 2.28 | 2.41 | 2 | 2.3429 | 2.31 | 2.3154 | 2.56 | 2.78 | 2.38 | 2.01 |  | 24
| 1 | 3 | 2.64 | 2.5 | 2.8 | 2.262 | 2.2228 | 2 | 2.8 | 2.6 | 2.3 | 2 | 2 | 2.45 | 2.09 | 2.32 | 2.33 | < 2 | 1.9 | 2.1 | 2.29 | 3.5 | 2 | 2.423 | 2.32 | 2.3629 | 2.46 | 2.77 | 2.41 | 1.89 |  | 16
| 1 | 4 | 2.62 | 2.4 | 3.62 | 2.8224 | 2.2535 | < 1 | 2.71 | 2.6 | 2.4 | 2 | 1.8 | 2.41 | 2.42 | 2.2 | 2.3 | 2 | 1.8 | 2 | 2.14 | 2.76 | 3 | 2.4266 | 2.44 | 2.2931 | 2.51 | 2.95 | 2.33 | 1.85 |  | 2.3352
| 1 | 5 | 2.62 | 2.4 | 3.46 | 2.2306 | 2.2948 | < 1 | 2.78 | 2.2 | 2.3 | 2 | 2.2 | 2.35 | 2.52 | 2.3 | 2.31 | < 2 | 1.9 | 2 | 2.21 | 9.93 | 3 | 2.2798 | 2.43 | 2.2386 | 2.72 | 2.6 | 2.32 | 1.98 |  | 44
| 1 | 6 | 2.66 | 2.4 | 2.65 | 2.311 | 2.1519 | < 1 | 2.67 | 2.4 | 2.3 | 2 | 1.6 | 2.31 | 2.49 | 2.24 | 2.35 | < 2 | 1.9 | 2 | 2.24 | 2.96 | 2 | 2.2349 | 2.36 | 2.2625 | 2.57 | 2.62 | 2.22 | 1.89
| Mean |  | 2.6383 | 2.4 | 2.8883 | 2.4364 | 2.2208 | 1.6667 | 2.7817 | 2.3833 | 2.3333 | 2 | 1.9 | 2.3867 | 2.3933 | 2.285 | 2.2967 | 2 | 1.8833 | 2.0333 | 2.235 | 4.095 | 2.3333 | 2.3376 | 2.3733 | 2.2943 | 2.5517 | 2.7533 | 2.3417 | 1.9533
| Median |  | 2.635 | 2.4 | 2.725 | 2.3433 | 2.2351 | 2 | 2.79 | 2.35 | 2.3 | 2 | 1.9 | 2.38 | 2.435 | 2.27 | 2.32 | 2 | 1.9 | 2 | 2.245 | 2.985 | 2 | 2.3307 | 2.37 | 2.2933 | 2.535 | 2.775 | 2.355 | 1.935
| Std Dev. |  | 0.018348 | 0.063246 | 0.53361 | 0.23395 | 0.057284 | 0.57735 | 0.081343 | 0.18348 | 0.05164 | 0 | 0.20976 | 0.066232 | 0.15578 | 0.089833 | 0.090259 | 0 | 0.040825 | 0.05164 | 0.054681 | 2.8806 | 0.5164 | 0.076799 | 0.054283 | 0.043051 | 0.092394 | 0.12894 | 0.069113 | 0.09395
| Rel.Std.Dev. |  | 0.6955% | 2.635% | 18.47% | 9.602% | 2.579% | 34.64% | 2.924% | 7.699% | 2.213% | 0% | 11.04% | 2.775% | 6.509% | 3.931% | 3.93% | 0% | 2.168% | 2.54% | 2.447% | 70.34% | 22.13% | 3.285% | 2.287% | 1.876% | 3.621% | 4.683% | 2.951% | 4.81%
| PDM3 |  | 12.98% | 2.775% | 23.69% | 4.332% | -4.901% | -28.63% | 19.12% | 2.061% | -0.08005% | -14.35% | -18.64% | 2.204% | 2.489% | -2.15% | -1.65% | -14.35% | -19.35% | -12.93% | -4.291% | 75.36% | -0.08005% | 0.1029% | 1.633% | -1.751% | 9.27% | 17.91% | 0.2768% | -16.35%
| Z-Score (Absolute) |  | 1.8 | 0.36 | 3.31 | 0.58 | 0.72 | 7.58 | 2.67 | 0.26 | 0.04 | NA | 2.65 | 0.28 | 0.32 | 0.33 | 0.26 | 6.07 | 2.75 | 1.85 | 0.63 | 10.59 | NA | 0.01 | 0.2 | 0.27 | 1.28 | 2.5 | 0.01 | 2.33
Results from laboratories 1.12 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 233b (Certified Value 1.92 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.99 | 1.9441 | 1.7236 | 1.7859 | 1.8862 | 1.95 | 2.129 | 1.91 | 1.8211 | 1.87 | 1.97 | 1.97 | 1.94 | 1.91 | 1.88 | 1.93 | 1.908 | 1.855 | 1.863 | 1.579 | 1.8476 | 1.97 | 1.97 | 2.383 | 1.94 | 2.4 | 2.05 | 1.96 |  | 1
| 1 | 2 | 1.96 | 1.9271 | 1.7284 | 1.7992 | 1.9124 | 1.85 | 2.122 | 1.86 | 1.8187 | 1.86 | 2.01 | 2.05 | 1.88 | 1.97 | 1.86 | 1.98 | 1.92 | 1.88 | 1.864 | 1.61 | 1.8537 | 1.96 | 1.95 | 2.459 | 1.96 | 2.33 | 2.08 | 1.95 |  | 0
| 1 | 3 | 1.97 | 1.9436 | 1.8059 | 1.8036 | 1.8926 | 1.91 | 2.151 | 1.89 | 1.8175 | 1.82 | 1.91 | 2.04 | 1.88 | 1.88 | 1.9 | 1.95 | 1.898 | 1.93 | 1.822 | 2.06 | 1.9293 | 1.955 | 1.95 | 2.468 | 1.93 | 2.38 | 2.06 | 1.93 |  | 16
| 1 | 4 | 1.94 | 1.8915 | 1.9347 | 1.7911 | 1.9158 | 1.93 | 2.107 | 1.89 | 1.8185 | 1.84 | 1.99 | 2.04 | 1.88 | 1.89 | 1.89 | 1.97 | 1.926 | 1.928 | 1.81 | 1.856 | 1.8869 | 1.97 | 1.94 | 2.456 | 1.99 | 2.32 | 2.06 | 1.91 |  | 1.9197
| 1 | 5 | 1.92 | 1.8692 | 1.9186 | 1.8073 | 1.9433 | 1.87 | 2.122 | 1.9 | 1.8174 | 1.83 | 1.98 | 2.05 | 1.91 | 1.87 | 1.85 | 2.04 | 1.908 | 1.934 | 1.951 | 2.65 | 1.8635 | 1.975 | 1.94 | 2.362 | 1.96 | 2.24 | 2.06 | 1.96 |  | 45
| 1 | 6 | 1.97 | 1.9149 | 1.9301 | 1.7761 | 1.8088 | 1.88 | 2.136 | 1.95 | 1.8261 | 1.82 | 1.95 | 2.04 | 1.9 | 2 | 1.9 | 1.99 | 1.924 | 1.861 | 1.842 | 2.05 | 1.9255 | 1.97 | 1.95 | 2.415 | 1.96 | 2.37 | 2.06 | 1.99
| Mean |  | 1.9583 | 1.9151 | 1.8402 | 1.7939 | 1.8932 | 1.8983 | 2.1278 | 1.9 | 1.8199 | 1.84 | 1.9683 | 2.0317 | 1.8983 | 1.92 | 1.88 | 1.9767 | 1.914 | 1.898 | 1.8587 | 1.9675 | 1.8844 | 1.9667 | 1.95 | 2.4238 | 1.9567 | 2.34 | 2.0617 | 1.95
| Median |  | 1.965 | 1.921 | 1.8622 | 1.7952 | 1.9025 | 1.895 | 2.1255 | 1.895 | 1.8186 | 1.835 | 1.975 | 2.04 | 1.89 | 1.9 | 1.885 | 1.975 | 1.914 | 1.904 | 1.8525 | 1.953 | 1.8752 | 1.97 | 1.95 | 2.4355 | 1.96 | 2.35 | 2.06 | 1.955
| Std Dev. |  | 0.024833 | 0.02987 | 0.10044 | 0.011751 | 0.045954 | 0.038166 | 0.014878 | 0.029665 | 0.0033253 | 0.020976 | 0.034881 | 0.030605 | 0.024014 | 0.052915 | 0.020976 | 0.037771 | 0.011027 | 0.036775 | 0.050127 | 0.39302 | 0.035895 | 0.0075277 | 0.010954 | 0.044251 | 0.020656 | 0.057619 | 0.0098319 | 0.027568
| Rel.Std.Dev. |  | 1.268% | 1.56% | 5.458% | 0.6551% | 2.427% | 2.011% | 0.6992% | 1.561% | 0.1827% | 1.14% | 1.772% | 1.506% | 1.265% | 2.756% | 1.116% | 1.911% | 0.5761% | 1.938% | 2.697% | 19.98% | 1.905% | 0.3828% | 0.5618% | 1.826% | 1.056% | 2.462% | 0.4769% | 1.414%
| PDM3 |  | 2.014% | -0.2399% | -4.14% | -6.553% | -1.38% | -1.112% | 10.84% | -1.025% | -5.198% | -4.15% | 2.535% | 5.834% | -1.112% | 0.01708% | -2.067% | 2.969% | -0.2955% | -1.129% | -3.178% | 2.491% | -1.838% | 2.448% | 1.58% | 26.26% | 1.927% | 21.9% | 7.397% | 1.58%
| Z-Score (Absolute) |  | 0.61 | 0.04 | 1.16 | 1.85 | 0.37 | 0.29 | 3.15 | 0.26 | 1.46 | 1.16 | 0.76 | 1.71 | 0.29 | 0.04 | 0.56 | 0.89 | 0.05 | 0.29 | 0.88 | 0.75 | 0.5 | 0.74 | 0.49 | 7.59 | 0.59 | 6.34 | 2.16 | 0.49
| Analytical results for Nb in OREAS 233b (Certified Value 3.54 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.8 | 3.65 | 19.55 | 2.2818 | 3.5674 | 2.1 | 4.2 | 3.6 | 3.4 | < 5 | 3.6 | 3.6 | 3.4 | 3.3 | 3.2 | 24 | 3.3 | 2.9 | 3.5 | 5 | 3.0343 | 3.5 | 5.2562 | 4.1 | 3.9 | 4 | 3.2 |  |  | 1
| 1 | 2 | 3.7 | 3.65 | 20.29 | 2.471 | 3.533 | 1.2 | 4 | 3.6 | 3.4 | < 5 | 3.6 | 3.7 | 3.3 | 3.6 | 3.4 | 24 | 3.1 | 2.9 | 3.6 | 9 | 2.9674 | 3.6 | 5.1596 | 4.1 | 3.7 | 3.9 | 3.1 |  |  | 7
| 1 | 3 | 3.7 | 3.79 | 19.42 | 2.2814 | 3.5114 | 2.6 | 4 | 3.6 | 3.44 | < 5 | 3.7 | 3.7 | 3.3 | 3.3 | 3.5 | 24 | 3.5 | 2.9 | 3.6 | 6 | 2.9126 | 3.6 | 5.2659 | 4 | 3.8 | 3.9 | 2.9 |  |  | 16
| 1 | 4 | 3.7 | 3.72 | 21.18 | 2.2611 | 3.5448 | 0.2 | 3.9 | 3.7 | 3.47 | < 5 | 3.7 | 3.6 | 3.4 | 3.5 | 3.4 | 24 | 3.1 | 3 | 3.6 | 6 | 3.0173 | 3.6 | 5.2205 | 4.1 | 3.8 | 4.1 | 2.9 |  |  | 3.5365
| 1 | 5 | 3.6 | 3.55 | 20.23 | 2.2389 | 3.6011 | 0.3 | 3.9 | 3.5 | 3.46 | < 5 | 3.7 | 3.6 | 3.5 | 3.5 | 3.5 | 25 | 3.4 | 3 | 3.6 | 6 | 2.9791 | 3.9 | 5.1144 | 4.3 | 3.6 | 4 | 3 |  |  | 46
| 1 | 6 | 3.5 | 3.67 | 18.33 | 2.2184 | 4.1128 | < 0.1 | 3.8 | 3.8 | 3.43 | < 5 | 3.7 | 3.6 | 3.6 | 3.3 | 3.5 | 24 | 3.3 | 3 | 3.4 | 5 | 3.1288 | 3.7 | 5.0781 | 4.2 | 3.7 | 4 | 2.9
| Mean |  | 3.6667 | 3.6717 | 19.8333 | 2.2921 | 3.6451 | 1.28 | 3.9667 | 3.6333 | 3.4333 |  | 3.6667 | 3.6333 | 3.4167 | 3.4167 | 3.4167 | 24.1667 | 3.2833 | 2.95 | 3.55 | 6.1667 | 3.0066 | 3.65 | 5.1825 | 4.1333 | 3.75 | 3.9833 | 3
| Median |  | 3.7 | 3.66 | 19.89 | 2.2713 | 3.5561 | 1.2 | 3.95 | 3.6 | 3.435 |  | 3.7 | 3.6 | 3.4 | 3.4 | 3.45 | 24 | 3.3 | 2.95 | 3.6 | 6 | 2.9982 | 3.6 | 5.1901 | 4.1 | 3.75 | 4 | 2.95
| Std Dev. |  | 0.10328 | 0.080104 | 0.96856 | 0.091051 | 0.23117 | 1.0663 | 0.13663 | 0.10328 | 0.029439 |  | 0.05164 | 0.05164 | 0.1169 | 0.13292 | 0.1169 | 0.40825 | 0.16021 | 0.054772 | 0.083666 | 1.472 | 0.073422 | 0.13784 | 0.077321 | 0.10328 | 0.10488 | 0.075277 | 0.12649
| Rel.Std.Dev. |  | 2.817% | 2.182% | 4.883% | 3.972% | 6.342% | 83.3% | 3.444% | 2.843% | 0.8575% |  | 1.408% | 1.421% | 3.422% | 3.89% | 3.422% | 1.689% | 4.879% | 1.857% | 2.357% | 23.87% | 2.442% | 3.776% | 1.492% | 2.499% | 2.797% | 1.89% | 4.216%
| PDM3 |  | 3.681% | 3.822% | 460.8% | -35.19% | 3.07% | -63.81% | 12.16% | 2.738% | -2.917% |  | 3.681% | 2.738% | -3.388% | -3.388% | -3.388% | 583.4% | -7.158% | -16.58% | 0.382% | 74.37% | -14.98% | 3.21% | 46.54% | 16.88% | 6.037% | 12.64% | -15.17%
| Z-Score (Absolute) |  | 0.07 | 0.08 | 32.77 | 2.71 | 0.02 | 5.18 | 0.67 | 0 | 0.4 | 2.29 | 0.07 | 0 | 0.44 | 0.44 | 0.44 | 41.54 | 0.71 | 1.38 | 0.17 | 5.12 | 1.27 | 0.03 | 3.13 | 1.01 | 0.24 | 0.71 | 1.28
| Analytical results for Nd in OREAS 233b (Certified Value 8.27 ppm) | [30 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*OES; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 8.55; 1.02: 4.25; 1.03: 7.9396; 1.04: 8.2; 1.05: 9.2; 1.07: 8; 1.08: 8.5; 1.10: 8.5092 |  |  | 1
| 1 | 2 | 1.01: 8.33; 1.02: 4.9; 1.03: 8.0588; 1.04: 8.1; 1.05: 8.6; 1.07: 7.8; 1.08: 8; 1.10: 8.5147 |  |  | 26
| 1 | 3 | 1.01: 8.25; 1.02: 4.81; 1.03: 7.9292; 1.04: 8.1; 1.05: 8.7; 1.07: 8; 1.08: 9; 1.10: 8.3784 |  |  | 16
| 1 | 4 | 1.01: 8.5; 1.02: 5.33; 1.03: 8.0879; 1.04: 7.7; 1.05: 8.8; 1.07: 8; 1.08: 8.5; 1.10: 8.4575 |  |  | 8.2715
| 1 | 5 | 1.01: 8.34; 1.02: 4.96; 1.03: 7.9067; 1.04: 7.6; 1.05: 8.9; 1.07: 8; 1.08: 8.5; 1.10: 8.1533 |  |  | 47
| 1 | 6 | 1.01: 8.55; 1.02: 4.52; 1.03: 7.9226; 1.04: 7.8; 1.05: 8.3; 1.07: 8; 1.08: 8; 1.10: 8.4216
| Mean |  | 1.01: 8.42; 1.02: 4.795; 1.03: 7.9741; 1.04: 7.9167; 1.05: 8.75; 1.07: 7.9667; 1.08: 8.4167; 1.10: 8.4058
| Median |  | 1.01: 8.42; 1.02: 4.855; 1.03: 7.9344; 1.04: 7.95; 1.05: 8.75; 1.07: 8; 1.08: 8.5; 1.10: 8.4395
| Std Dev. |  | 1.01: 0.12931; 1.02: 0.37346; 1.03: 0.078129; 1.04: 0.24833; 1.05: 0.30166; 1.07: 0.08165; 1.08: 0.37639; 1.10: 0.13414
| Rel.Std.Dev. |  | 1.01: 1.536%; 1.02: 7.788%; 1.03: 0.9798%; 1.04: 3.137%; 1.05: 3.448%; 1.07: 1.025%; 1.08: 4.472%; 1.10: 1.596%
| PDM3 |  | 1.01: 1.795%; 1.02: -42.03%; 1.03: -3.595%; 1.04: -4.29%; 1.05: 5.785%; 1.07: -3.685%; 1.08: 1.755%; 1.10: 1.624%
| Z-Score (Absolute) |  | 1.01: 0.68; 1.02: 10.02; 1.03: 0.64; 1.04: 0.81; 1.05: 1.65; 1.07: 0.66; 1.08: 0.67; 1.10: 0.64
| Analytical results for Ni in OREAS 233b (Certified Value 97 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 98.2 | 102.5 | 89.53 | 96.81 | 101.8294 | 95 | 124 | 99 | 97.4 | 90 | 96.6 | 104 | 101 | 95.7 | 95.1 | 88 | 97 | 94.9 | 94 | 80.25 | 90.5076 | 95 | 98.4036 | 99.6 | 118.5227 | 91.8 | 109.1 | 102.4 | 103 | 1
| 1 | 2 | 97.9 | 99.8 | 88.7 | 97.74 | 104.2841 | 96 | 117 | 96 | 99.4 | 88 | 96.8 | 109.5 | 97.2 | 102.5 | 97.7 | 90 | 97 | 96.9 | 95 | 78.15 | 91.419 | 95 | 103.2587 | 99.4 | 97.849 | 91.2 | 105 | 103.6 | 104 | 27
| 1 | 3 | 97.9 | 100.5 | 95.33 | 96.96 | 101.8498 | 95 | 120 | 98 | 98 | 88 | 95.4 | 108.5 | 101 | 97.6 | 99.6 | 89 | 95 | 95.5 | 94 | 79.56 | 94.9623 | 95 | 100.8186 | 99.9 | 97.8294 | 91.3 | 104.2 | 102.2 | 99.5 | 16
| 1 | 4 | 96.7 | 99.6 | 93.79 | 98.31 | 103.1788 | 96 | 117 | 99 | 98.6 | 91 | 96.5 | 108 | 97.7 | 96.4 | 100 | 89 | 98 | 95.4 | 92 | 77.99 | 94.9437 | 95 | 101.0472 | 99.2 | 92.1324 | 90.1 | 105.7 | 102.8 | 100 | 97.3957
| 1 | 5 | 95.2 | 98.6 | 92.86 | 97.87 | 105.9614 | 96 | 115 | 98 | 96.7 | 89 | 95.4 | 106 | 100 | 98.8 | 99.5 | 91 | 95 | 95.8 | 99 | 81.63 | 93.93 | 95 | 97.2875 | 101.5 | 88.7218 | 91 | 104.6 | 102.5 | 102 | 48
| 1 | 6 | 98 | 101.4 | 92.74 | 97.23 | 97.7 | 95 | 114 | 100 | 98 | 89 | 94.6 | 107.5 | 99.6 | 99.7 | 101 | 88 | 97 | 96.5 | 93 | 79.19 | 94.9623 | 95 | 101.2374 | 101.5 | 95.514 | 88.8 | 103.7 | 101.7 | 103
| Mean |  | 97.3167 | 100.4 | 92.1583 | 97.4867 | 102.4673 | 95.5 | 117.8333 | 98.3333 | 98.0167 | 89.1667 | 95.8833 | 107.25 | 99.4167 | 98.45 | 98.8167 | 89.1667 | 96.5 | 95.8333 | 94.5 | 79.4617 | 93.4541 | 95 | 100.3422 | 100.1833 | 98.4282 | 90.7 | 105.3833 | 102.5333 | 101.9167
| Median |  | 97.9 | 100.15 | 92.8 | 97.485 | 102.5143 | 95.5 | 117 | 98.5 | 98 | 89 | 95.95 | 107.75 | 99.8 | 98.2 | 99.55 | 89 | 97 | 95.65 | 94 | 79.375 | 94.4369 | 95 | 100.9329 | 99.75 | 96.6717 | 91.1 | 104.8 | 102.45 | 102.5
| Std Dev. |  | 1.1652 | 1.39 | 2.5464 | 0.5814 | 2.8118 | 0.54772 | 3.656 | 1.3663 | 0.9347 | 1.169 | 0.87731 | 1.9685 | 1.6278 | 2.4729 | 2.1123 | 1.169 | 1.2247 | 0.74207 | 2.429 | 1.3633 | 1.9909 | 0 | 2.1504 | 1.0458 | 10.4578 | 1.0844 | 1.9447 | 0.6377 | 1.8005
| Rel.Std.Dev. |  | 1.197% | 1.384% | 2.763% | 0.5964% | 2.744% | 0.5735% | 3.103% | 1.389% | 0.9536% | 1.311% | 0.915% | 1.835% | 1.637% | 2.512% | 2.138% | 1.311% | 1.269% | 0.7743% | 2.57% | 1.716% | 2.13% | 0% | 2.143% | 1.044% | 10.62% | 1.196% | 1.845% | 0.6219% | 1.767%
| PDM3 |  | -0.08117% | 3.085% | -5.377% | 0.09337% | 5.207% | -1.946% | 20.98% | 0.9627% | 0.6375% | -8.449% | -1.553% | 10.12% | 2.075% | 1.082% | 1.459% | -8.449% | -0.9197% | -1.604% | -2.973% | -18.41% | -4.047% | -2.46% | 3.025% | 2.862% | 1.06% | -6.875% | 8.201% | 5.275% | 4.642%
| Z-Score (Absolute) |  | 0.19 | 0.64 | 1.57 | 0.14 | 1.19 | 0.67 | 5.31 | 0.08 | 0 | 2.37 | 0.57 | 2.47 | 0.38 | 0.12 | 0.21 | 2.37 | 0.41 | 0.58 | 0.94 | 4.97 | 1.22 | 0.81 | 0.62 | 0.58 | 0.11 | 1.96 | 1.97 | 1.21 | 1.04
| Analytical results for P in OREAS 233b (Certified Value 0.043 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.0429 | 0.0449 | 0.041914 | 0.041644 | 0.039 | 0.0502 | 0.046 | 0.0413 | 0.0424 | 0.042 | 0.045 | 0.045 | 0.043 | 0.042 | 0.0413 | 0.044 | 0.042 | 0.043 | 0.042 | 0.043 | 0.040023 | 0.0435 | 0.0443 | 0.0428 | 0.041 |  |  |  |  | 1
| 1 | 2 | 0.0428 | 0.0438 | 0.04211 | 0.042555 | 0.038 | 0.0528 | 0.045 | 0.0445 | 0.0426 | 0.043 | 0.046 | 0.043 | 0.042 | 0.043 | 0.0423 | 0.044 | 0.041 | 0.043 | 0.042 | 0.045 | 0.042127 | 0.0431 | 0.0428 | 0.043 | 0.0404 |  |  |  |  | 0
| 1 | 3 | 0.0426 | 0.0448 | 0.041377 | 0.041678 | 0.04 | 0.0515 | 0.046 | 0.0445 | 0.0431 | 0.042 | 0.046 | 0.044 | 0.042 | 0.043 | 0.042 | 0.043 | 0.042 | 0.043 | 0.04 | 0.044 | 0.041959 | 0.0428 | 0.0433 | 0.0426 | 0.0399 |  |  |  |  | 16
| 1 | 4 | 0.0426 | 0.0433 | 0.041844 | 0.042386 | 0.037 | 0.0506 | 0.046 | 0.0431 | 0.0431 | 0.042 | 0.046 | 0.043 | 0.044 | 0.043 | 0.0417 | 0.044 | 0.041 | 0.042 | 0.0415 | 0.044 | 0.041849 | 0.0423 | 0.0426 | 0.0428 | 0.0401 |  |  |  |  | 0.042905
| 1 | 5 | 0.0418 | 0.045 | 0.041721 | 0.043196 | 0.038 | 0.0506 | 0.045 | 0.0397 | 0.0432 | 0.042 | 0.046 | 0.045 | 0.043 | 0.043 | 0.0429 | 0.044 | 0.041 | 0.045 | 0.0425 | 0.044 | 0.041322 | 0.0438 | 0.0421 | 0.0425 | 0.0406 |  |  |  |  | 49
| 1 | 6 | 0.0428 | 0.045 | 0.042154 | 0.039911 | 0.035 | 0.0515 | 0.046 | 0.0437 | 0.0426 | 0.041 | 0.046 | 0.042 | 0.047 | 0.043 | 0.0415 | 0.045 | 0.042 | 0.043 | 0.042 | 0.045 | 0.040728 | 0.0424 | 0.0434 | 0.0424 | 0.0409
| Mean |  | 0.042583 | 0.044467 | 0.041853 | 0.041895 | 0.037833 | 0.0512 | 0.045667 | 0.0428 | 0.042833 | 0.042 | 0.045833 | 0.043667 | 0.0435 | 0.042833 | 0.04195 | 0.044 | 0.0415 | 0.043167 | 0.041667 | 0.044167 | 0.041335 | 0.042983 | 0.043083 | 0.042683 | 0.040483
| Median |  | 0.0427 | 0.04485 | 0.041879 | 0.042032 | 0.038 | 0.05105 | 0.046 | 0.0434 | 0.04285 | 0.042 | 0.046 | 0.0435 | 0.043 | 0.043 | 0.04185 | 0.044 | 0.0415 | 0.043 | 0.042 | 0.044 | 0.041586 | 0.04295 | 0.04305 | 0.0427 | 0.0405
| Std Dev. |  | 0.00040208 | 0.00073121 | 0.00028441 | 0.0011328 | 0.0017224 | 0.00094446 | 0.0005164 | 0.0019256 | 0.00033862 | 0.00063246 | 0.00040825 | 0.0012111 | 0.0018708 | 0.00040825 | 0.00058566 | 0.00063246 | 0.00054772 | 0.00098319 | 0.0008756 | 0.00075277 | 0.00082095 | 0.00059805 | 0.00076267 | 0.00022286 | 0.00043551
| Rel.Std.Dev. |  | 0.9442% | 1.644% | 0.6795% | 2.704% | 4.553% | 1.845% | 1.131% | 4.499% | 0.7906% | 1.506% | 0.8907% | 2.773% | 4.301% | 0.9531% | 1.396% | 1.437% | 1.32% | 2.278% | 2.101% | 1.704% | 1.986% | 1.391% | 1.77% | 0.5221% | 1.076%
| PDM3 |  | -0.7499% | 3.64% | -2.451% | -2.355% | -11.82% | 19.33% | 6.437% | -0.2449% | -0.1672% | -2.109% | 6.825% | 1.775% | 1.387% | -0.1672% | -2.226% | 2.552% | -3.275% | 0.6097% | -2.886% | 2.94% | -3.66% | 0.1824% | 0.4155% | -0.5168% | -5.644%
| Z-Score (Absolute) |  | 0.18 | 1.17 | 0.7 | 0.67 | 3.59 | 6.01 | 2.04 | 0.02 | 0 | 0.6 | 2.16 | 0.6 | 0.48 | 0 | 0.63 | 0.84 | 0.96 | 0.24 | 0.84 | 0.96 | 1.08 | 0.11 | 0.18 | 0.11 | 1.69
| Analytical results for Pb in OREAS 233b (Certified Value 23.9 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 22.7 | 25 | 17.25 | 22.019 | 24.5811 | 18 | 29.1 | 26 | 23.6 | 25 | 20.1 | 23.9 | 25.2 | 23 | 24 | 17 | 23 | 23.1 | 26.1 | 26 | 21.6294 | 25 | 24.0788 | 24.1 | 88.4449 | 24.3 | 25.7 | 25.3 | 22.9 | 1
| 1 | 2 | 22.7 | 24.7 | 19.89 | 22.4573 | 25.5544 | 18 | 27.6 | 25 | 23.8 | 25 | 20.7 | 24.6 | 24.8 | 25.3 | 23.8 | 17 | 23 | 22.6 | 25.8 | 24.45 | 22.3626 | 24 | 23.8898 | 24.3 | 89.1399 | 24.3 | 25.1 | 24.9 | 22.4 | 28
| 1 | 3 | 22.7 | 25 | 20.43 | 22.0393 | 24.8503 | 19 | 27.9 | 26 | 23.8 | 26 | 21.1 | 24.9 | 22.9 | 23.5 | 25.2 | 18 | 23 | 22.6 | 25.6 | 23.53 | 22.4378 | 23 | 24.1009 | 23.9 | 88.2201 | 24.3 | 24.6 | 26 | 20.4 | 16
| 1 | 4 | 22.3 | 25 | 20.61 | 21.7862 | 24.7539 | 19 | 28 | 25 | 24.2 | 25 | 21.5 | 24.3 | 24.8 | 22.5 | 25.3 | 18 | 23 | 23.2 | 26.2 | 24.24 | 21.6106 | 24 | 24.7695 | 24.2 | 87.2986 | 24.3 | 24.7 | 26.4 | 20.4 | 23.9222
| 1 | 5 | 22.1 | 24.6 | 20.97 | 22.527 | 25.5226 | 19 | 27.8 | 25 | 24.2 | 25 | 20.2 | 24.5 | 26.2 | 22.6 | 25.2 | 17 | 23 | 23.1 | 25.6 | 30.36 | 20.9808 | 24 | 23.4906 | 24.8 | 84.1799 | 24.2 | 24.8 | 26.9 | 21.2 | 50
| 1 | 6 | 22.6 | 25.7 | 22.06 | 22.2518 | 24.018 | 19 | 26.8 | 26 | 23.9 | 25 | 20.5 | 23.9 | 24.6 | 22.5 | 24.4 | 17 | 23 | 22.7 | 25.5 | 25.54 | 21.902 | 24 | 24.1002 | 24 | 87.8795 | 24 | 24.6 | 26.5 | 20.9
| Mean |  | 22.5167 | 25 | 20.2017 | 22.1801 | 24.88 | 18.6667 | 27.8667 | 25.5 | 23.9167 | 25.1667 | 20.6833 | 24.35 | 24.75 | 23.2333 | 24.65 | 17.3333 | 23 | 22.8833 | 25.8 | 25.6867 | 21.8205 | 24 | 24.0716 | 24.2167 | 87.5272 | 24.2333 | 24.9167 | 26 | 21.3667
| Median |  | 22.65 | 25 | 20.52 | 22.1456 | 24.8021 | 19 | 27.85 | 25.5 | 23.85 | 25 | 20.6 | 24.4 | 24.8 | 22.8 | 24.8 | 17 | 23 | 22.9 | 25.7 | 24.995 | 21.7657 | 24 | 24.0895 | 24.15 | 88.0498 | 24.3 | 24.75 | 26.2 | 21.05
| Std Dev. |  | 0.25626 | 0.38471 | 1.6169 | 0.28399 | 0.586 | 0.5164 | 0.74207 | 0.54772 | 0.24014 | 0.40825 | 0.53821 | 0.39875 | 1.0728 | 1.0838 | 0.66858 | 0.5164 | 0 | 0.27869 | 0.28983 | 2.459 | 0.54171 | 0.63246 | 0.41458 | 0.31885 | 1.7495 | 0.12111 | 0.42622 | 0.7642 | 1.052
| Rel.Std.Dev. |  | 1.138% | 1.539% | 8.004% | 1.28% | 2.355% | 2.766% | 2.663% | 2.148% | 1.004% | 1.622% | 2.602% | 1.638% | 4.335% | 4.665% | 2.712% | 2.979% | 0% | 1.218% | 1.123% | 9.573% | 2.483% | 2.635% | 1.722% | 1.317% | 1.999% | 0.4997% | 1.711% | 2.939% | 4.923%
| PDM3 |  | -5.875% | 4.506% | -15.55% | -7.282% | 4.004% | -21.97% | 16.49% | 6.596% | -0.02299% | 5.202% | -13.54% | 1.788% | 3.461% | -2.879% | 3.043% | -27.54% | -3.855% | -4.343% | 7.85% | 7.376% | -8.785% | 0.3254% | 0.6248% | 1.231% | 265.9% | 1.301% | 4.157% | 8.686% | -10.68%
| Z-Score (Absolute) |  | 0.89 | 0.41 | 2.11 | 1.07 | 0.35 | 2.92 | 1.92 | 0.67 | 0.16 | 0.5 | 1.86 | 0.07 | 0.28 | 0.52 | 0.23 | 3.62 | 0.64 | 0.7 | 0.83 | 0.77 | 1.26 | 0.11 | 0.08 | 0 | 33.27 | 0.01 | 0.37 | 0.94 | 1.5
| Analytical results for Pr in OREAS 233b (Certified Value 1.71 ppm) | [30 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*OES; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS | 2
3
| 1 | 1 | 1.01: 1.85; 1.02: 1.49; 1.03: 1.6218; 1.04: 1.8; 1.05: 2.1; 1.07: 1.75; 1.08: 1.7; 1.10: 1.731 | 1
| 1 | 2 | 1.01: 1.78; 1.02: 1.63; 1.03: 1.616; 1.04: 1.7; 1.05: 2; 1.07: 1.72; 1.08: 1.9; 1.10: 1.7912 | 29
| 1 | 3 | 1.01: 1.78; 1.02: 1.59; 1.03: 1.6305; 1.04: 1.7; 1.05: 2; 1.07: 1.79; 1.08: 1.8; 1.10: 1.7536 | 16
| 1 | 4 | 1.01: 1.84; 1.02: 1.68; 1.03: 1.6042; 1.04: 1.7; 1.05: 2; 1.07: 1.78; 1.08: 1.7; 1.10: 1.7642 | 1.7099
| 1 | 5 | 1.01: 1.8; 1.02: 1.62; 1.03: 1.6516; 1.04: 1.7; 1.05: 2.1; 1.07: 1.71; 1.08: 1.7; 1.10: 1.7133 | 51
| 1 | 6 | 1.01: 1.79; 1.02: 1.54; 1.03: 1.6076; 1.04: 1.7; 1.05: 1.9; 1.07: 1.69; 1.08: 1.6; 1.10: 1.8013
| Mean |  | 1.01: 1.8067; 1.02: 1.5917; 1.03: 1.622; 1.04: 1.7167; 1.05: 2.0167; 1.07: 1.74; 1.08: 1.7333; 1.10: 1.7591
| Median |  | 1.01: 1.795; 1.02: 1.605; 1.03: 1.6189; 1.04: 1.7; 1.05: 2; 1.07: 1.735; 1.08: 1.7; 1.10: 1.7589
| Std Dev. |  | 1.01: 0.030768; 1.02: 0.067946; 1.03: 0.017368; 1.04: 0.040825; 1.05: 0.075277; 1.07: 0.04; 1.08: 0.10328; 1.10: 0.033925
| Rel.Std.Dev. |  | 1.01: 1.703%; 1.02: 4.269%; 1.03: 1.071%; 1.04: 2.378%; 1.05: 3.733%; 1.07: 2.299%; 1.08: 5.958%; 1.10: 1.929%
| PDM3 |  | 1.01: 5.659%; 1.02: -6.915%; 1.03: -5.144%; 1.04: 0.3951%; 1.05: 17.94%; 1.07: 1.76%; 1.08: 1.37%; 1.10: 2.876%
| Z-Score (Absolute) |  | 1.01: 1.02; 1.02: 2.12; 1.03: 1.67; 1.04: 0.29; 1.05: 4.09; 1.07: 0.05; 1.08: 0.05; 1.10: 0.33
| Analytical results for Rb in OREAS 233b (Certified Value 14 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 14.5 | 13.9 | 12.8011 | 12.912 | 14.4 | 15.5 | 14.5 | 13.5 | 12.6 | 15 | 13.8 | 13 | 13.3 | 14 | 13.9 | 14 | 14.5 | 14.3672 | 14 | 15.1473 | 15.1 | 14.5 | 15.7 | 12.8 | [5 empty] | 1
| 1 | 2 | 14.8 | 13.62 | 13.1725 | 13.0507 | 14.1 | 14.4 | 13.9 | 14.19 | 13.5 | 15.6 | 13.2 | 14 | 13.6 | 13.4 | 13.8 | 14.2 | 14 | 15.5209 | 14.3 | 15.0729 | 15.6 | 14.1 | 15.3 | 12.1 | [5 empty] | 30
| 1 | 3 | 14.6 | 13.81 | 13.0828 | 12.8028 | 14.1 | 14.6 | 14.1 | 13.68 | 11.5 | 15.7 | 13.4 | 13.3 | 13.9 | 14.3 | 14 | 14.8 | 14 | 15.909 | 14.2 | 15.3224 | 15.4 | 14.2 | 15.4 | 11.6 | [5 empty] | 16
| 1 | 4 | 14.3 | 13.77 | 13.1345 | 13.0173 | 13.7 | 14.4 | 14.2 | 14.08 | 13.7 | 15.2 | 14.2 | 13 | 14.1 | 13.3 | 13.6 | 14.2 | 15 | 14.6427 | 13.4 | 15.0154 | 15.9 | 14.1 | 16.1 | 11.6 | [5 empty] | 14.0247
| 1 | 5 | 14.4 | 13.6 | 13.2537 | 13.0611 | 13.8 | 14.5 | 13.9 | 13.85 | 14.7 | 15.3 | 14 | 13.2 | 14.1 | 14.1 | 13.9 | 14.1 | 14.5 | 14.0065 | 15.4 | 14.9532 | 15.5 | 13.8 | 16 | 11.9 | [5 empty] | 52
| 1 | 6 | 14.1 | 13.66 | 12.7618 | 12.7699 | 13.8 | 13.8 | 14.6 | 13.91 | 14 | 15.2 | 12.5 | 13.2 | 14.2 | 13.9 | 13.9 | 14.3 | 14 | 14.5436 | 14.4 | 14.9827 | 15.4 | 13.6 | 15.9 | 11.4
| Mean |  | 14.45 | 13.7267 | 13.0344 | 12.9356 | 13.9833 | 14.5333 | 14.2 | 13.8683 | 13.3333 | 15.3333 | 13.5167 | 13.2833 | 13.8667 | 13.8333 | 13.85 | 14.2667 | 14.3333 | 14.8317 | 14.2833 | 15.0823 | 15.4833 | 14.05 | 15.7333 | 11.9
| Median |  | 14.45 | 13.715 | 13.1087 | 12.9646 | 13.95 | 14.45 | 14.15 | 13.88 | 13.6 | 15.25 | 13.6 | 13.2 | 14 | 13.95 | 13.9 | 14.2 | 14.25 | 14.5931 | 14.25 | 15.0442 | 15.45 | 14.1 | 15.8 | 11.75
| Std Dev. |  | 0.2429 | 0.11894 | 0.20408 | 0.12754 | 0.26394 | 0.55015 | 0.29665 | 0.25357 | 1.129 | 0.26583 | 0.62102 | 0.37103 | 0.35024 | 0.39833 | 0.13784 | 0.28048 | 0.40825 | 0.72815 | 0.65243 | 0.1364 | 0.26394 | 0.31464 | 0.3266 | 0.50596
| Rel.Std.Dev. |  | 1.681% | 0.8665% | 1.566% | 0.986% | 1.888% | 3.785% | 2.089% | 1.828% | 8.468% | 1.734% | 4.594% | 2.793% | 2.526% | 2.879% | 0.9952% | 1.966% | 2.848% | 4.909% | 4.568% | 0.9044% | 1.705% | 2.239% | 2.076% | 4.252%
| PDM3 |  | 3.032% | -2.125% | -7.061% | -7.765% | -0.2951% | 3.627% | 1.25% | -1.115% | -4.93% | 9.331% | -3.623% | -5.286% | -1.127% | -1.365% | -1.246% | 1.725% | 2.201% | 5.754% | 1.844% | 7.541% | 10.4% | 0.1803% | 12.18% | -15.15%
| Z-Score (Absolute) |  | 0.63 | 0.42 | 1.42 | 1.56 | 0.05 | 0.75 | 0.26 | 0.21 | 0.99 | 1.9 | 0.72 | 1.06 | 0.22 | 0.26 | 0.24 | 0.36 | 0.46 | 1.18 | 0.39 | 1.54 | 2.12 | 0.05 | 2.48 | 3.06
| Analytical results for Re in OREAS 233b (Certified Value 0.003 ppm) | [30 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.29 | 1.30 | 1.31 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.002 | 0.002 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0196 | 0.004 | 0.004 | 0.003 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.003 | 0.005 | 0.004 | 0.002 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0174 | 0.004 | 0.003 | 0.004 | [10 empty] | 31
| 1 | 3 | < 0.002 | < 0.01 | 0.003 | < 0.05 | 0.004 | < 0.005 | 0.002 | 0.002 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0196 | 0.005 | 0.004 | 0.004 | [10 empty] | 16
| 1 | 4 | 0.002 | < 0.01 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0154 | 0.005 | 0.003 | 0.004 | [10 empty] | 0.0031
| 1 | 5 | < 0.002 | < 0.01 | 0.002 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.002 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.003 | 0.0161 | 0.005 | 0.003 | 0.004 | [10 empty] | 53
| 1 | 6 | 0.004 | < 0.01 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0165 | 0.003 | 0.003 | 0.004
| Mean |  | 1.01: 0.003; 1.03: 0.0021667; 1.05: 0.0032; 1.07: 0.005; 1.08: 0.0026667; 1.10: 0.002; 1.11: 0.0026667; 1.12: 0.0028333; 1.18: 0.0021667; 1.19: 0.017433; 1.22: 0.0043333; 1.23: 0.0033333; 1.24: 0.0038333
| Median |  | 1.01: 0.003; 1.03: 0.002; 1.05: 0.003; 1.07: 0.005; 1.08: 0.0025; 1.10: 0.002; 1.11: 0.003; 1.12: 0.003; 1.18: 0.002; 1.19: 0.01695; 1.22: 0.0045; 1.23: 0.003; 1.24: 0.004
| Std Dev. |  | 1.01: 0.0014142; 1.03: 0.00040825; 1.05: 0.00044721; 1.08: 0.0008165; 1.10: 0; 1.11: 0.0005164; 1.12: 0.00040825; 1.18: 0.00040825; 1.19: 0.0017985; 1.22: 0.0008165; 1.23: 0.0005164; 1.24: 0.00040825
| Rel.Std.Dev. |  | 1.01: 47.14%; 1.03: 18.84%; 1.05: 13.98%; 1.08: 30.62%; 1.10: 0%; 1.11: 19.36%; 1.12: 14.41%; 1.18: 18.84%; 1.19: 10.32%; 1.22: 18.84%; 1.23: 15.49%; 1.24: 10.65%
| PDM3 |  | 1.01: -3.226%; 1.03: -30.11%; 1.05: 3.226%; 1.07: 61.29%; 1.08: -13.98%; 1.10: -35.48%; 1.11: -13.98%; 1.12: -8.602%; 1.18: -30.11%; 1.19: 462.4%; 1.22: 39.78%; 1.23: 7.527%; 1.24: 23.66%
| Z-Score (Absolute) |  | 0.92 | 1.53 | 0.55 | 16.24 | 0.06 | 0 | 0.18 | 0.67 | 0.18 | 0.06 | 1836.88 | 16.24 | 0.31 | 16.24 | 0.55 | 10.68 | 1.04 | 0.31 | 0.67
| Analytical results for S in OREAS 233b (Certified Value 0.373 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.38 | 0.39 | 0.38596 | 0.3678 | 0.35 | 0.427 | 0.376 | 0.34 | 0.38 | 0.36 | 0.38 | 0.38 | 0.36 | 0.37 | 0.35 | 0.381 | 0.3 | 0.384 | 0.33932 | 0.375 | 0.37 | 0.32127 | 0.38 | 0.39 | 0.38 | 0.39 |  |  |  | 1
| 1 | 2 | 0.39 | 0.39 | 0.38747 | 0.37242 | 0.36 | 0.431 | 0.368 | 0.33 | 0.37 | 0.36 | 0.39 | 0.37 | 0.37 | 0.37 | 0.36 | 0.381 | 0.3 | 0.39 | 0.34447 | 0.375 | 0.37 | 0.34078 | 0.4 | 0.38 | 0.38 | 0.38 |  |  |  | 16
| 1 | 3 | 0.38 | 0.38 | 0.38442 | 0.37436 | 0.36 | 0.427 | 0.376 | 0.33 | 0.37 | 0.35 | 0.39 | 0.37 | 0.35 | 0.37 | 0.35 | 0.383 | 0.3 | 0.388 | 0.35569 | 0.375 | 0.36 | 0.33066 | 0.38 | 0.38 | 0.38 | 0.37 |  |  |  | 16
| 1 | 4 | 0.38 | 0.38 | 0.38713 | 0.38362 | 0.35 | 0.423 | 0.38 | 0.37 | 0.37 | 0.35 | 0.39 | 0.37 | 0.36 | 0.37 | 0.36 | 0.386 | 0.3 | 0.38 | 0.35402 | 0.375 | 0.36 | 0.33389 | 0.4 | 0.38 | 0.38 | 0.38 |  |  |  | 0.37327
| 1 | 5 | 0.38 | 0.38 | 0.38341 | 0.38899 | 0.35 | 0.436 | 0.376 | 0.37 | 0.38 | 0.36 | 0.39 | 0.38 | 0.36 | 0.36 | 0.37 | 0.379 | 0.3 | 0.408 | 0.3491 | 0.375 | 0.37 | 0.32712 | 0.38 | 0.37 | 0.38 | 0.38 |  |  |  | 54
| 1 | 6 | 0.39 | 0.38 | 0.38098 | 0.36948 | 0.34 | 0.414 | 0.382 | 0.37 | 0.38 | 0.34 | 0.39 | 0.37 | 0.39 | 0.37 | 0.36 | 0.384 | 0.3 | 0.382 | 0.35616 | 0.375 | 0.37 | 0.34657 | 0.4 | 0.38 | 0.38 | 0.39
| Mean |  | 0.38333 | 0.38333 | 0.3849 | 0.37611 | 0.35167 | 0.42633 | 0.37633 | 0.35167 | 0.375 | 0.35333 | 0.38833 | 0.37333 | 0.365 | 0.36833 | 0.35833 | 0.38233 | 0.3 | 0.38867 | 0.34979 | 0.375 | 0.36667 | 0.33338 | 0.39 | 0.38 | 0.38 | 0.38167
| Median |  | 0.38 | 0.38 | 0.38519 | 0.37339 | 0.35 | 0.427 | 0.376 | 0.355 | 0.375 | 0.355 | 0.39 | 0.37 | 0.36 | 0.37 | 0.36 | 0.382 | 0.3 | 0.386 | 0.35156 | 0.375 | 0.37 | 0.33227 | 0.39 | 0.38 | 0.38 | 0.38
| Std Dev. |  | 0.005164 | 0.005164 | 0.0024705 | 0.0083897 | 0.0075277 | 0.0074744 | 0.0048028 | 0.020412 | 0.0054772 | 0.008165 | 0.0040825 | 0.005164 | 0.013784 | 0.0040825 | 0.0075277 | 0.0025033 | 0 | 0.010172 | 0.006808 | 0 | 0.005164 | 0.00919 | 0.010954 | 0.0063246 | 6.0809e-17 | 0.0075277
| Rel.Std.Dev. |  | 1.347% | 1.347% | 0.6419% | 2.231% | 2.141% | 1.753% | 1.276% | 5.804% | 1.461% | 2.311% | 1.051% | 1.383% | 3.776% | 1.108% | 2.101% | 0.6548% | 0% | 2.617% | 1.946% | 0% | 1.408% | 2.757% | 2.809% | 1.664% | 1.6e-14% | 1.972%
| PDM3 |  | 2.697% | 2.697% | 3.116% | 0.7629% | -5.786% | 14.22% | 0.8219% | -5.786% | 0.4647% | -5.34% | 4.037% | 0.01822% | -2.214% | -1.321% | -4% | 2.429% | -19.63% | 4.126% | -6.288% | 0.4647% | -1.768% | -10.69% | 4.483% | 1.804% | 1.804% | 2.251%
| Z-Score (Absolute) |  | 0.55 | 0.55 | 0.67 | 0 | 1.88 | 3.86 | 0.02 | 1.88 | 0.09 | 1.75 | 0.94 | 0.21 | 0.85 | 0.6 | 1.36 | 0.48 | NA | 0.96 | 2.02 | 0.09 | 0.73 | 3.28 | 1.07 | 0.3 | 0.3 | 0.43
Results from laboratory 1.23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 233b (Certified Value 1.31 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.46 | 1.31 | 1.61 | 1 | 1.3 | 1.34 | < 5 | 1.7 | 1.29 | 1.38 | 1.22 | 1.25 | < 5 | 1.2 | 1.2 | 1.19 | 4.5 | 1.5 | 1.1017 | 1.29 | 1.3633 | 1.25 | 1.39 | 1.5 | 1.24 |  |  |  |  | 1
| 1 | 2 | 1.48 | 1.34 | 1.3491 | 0.9 | 1.3 | 1.28 | < 5 | 1.4 | 1.35 | 1.38 | 1.43 | 1.35 | < 5 | 1.2 | 1.2 | 1.25 | 5.6 | 1.5 | 1.0817 | 1.28 | 1.316 | 1.27 | 1.34 | 1.4 | 1.17 |  |  |  |  | 32
| 1 | 3 | 1.47 | 1.36 | 1.3055 | 1.1 | 1.3 | 1.33 | < 5 | 1.4 | 1.36 | 1.23 | 1.22 | 1.35 | < 5 | 1.2 | 1.2 | 1.25 | 7.82 | 1 | 1.0507 | 1.22 | 1.3317 | 1.27 | 1.33 | 1.4 | 1.13 |  |  |  |  | 16
| 1 | 4 | 1.44 | 1.36 | 1.3752 | 0.1 | 1.3 | 1.29 | < 5 | 1.6 | 1.33 | 1.4 | 1.29 | 1.4 | < 5 | 1.1 | 1.1 | 1.24 | 5.92 | 1.5 | 1.1316 | 1.22 | 1.3139 | 1.3 | 1.34 | 1.4 | 1.09 |  |  |  |  | 1.3066
| 1 | 5 | 1.39 | 1.28 | 1.372 | < 0.1 | 1.1 | 1.31 | < 5 | 1.6 | 1.33 | 1.43 | 1.28 | 1.38 | < 5 | 1.2 | 1.2 | 1.28 | 4.83 | 1.5 | 1.0711 | 1.28 | 1.2691 | 1.5 | 1.37 | 1.5 | 1.16 |  |  |  |  | 55
| 1 | 6 | 1.48 | 1.33 | 1.4139 | 0.4 | 1.3 | 1.26 | < 5 | 1.7 | 1.34 | 1.3 | 1.23 | 1.3 | < 5 | 1.2 | 1.2 | 1.22 | 5.17 | 1 | 1.1381 | 1.33 | 1.301 | 1.31 | 1.32 | 1.5 | 1.09
| Mean |  | 1.4533 | 1.33 | 1.4043 | 0.7 | 1.2667 | 1.3017 |  | 1.5667 | 1.3333 | 1.3533 | 1.2783 | 1.3383 |  | 1.1833 | 1.1833 | 1.2383 | 5.64 | 1.3333 | 1.0958 | 1.27 | 1.3158 | 1.3167 | 1.3483 | 1.45 | 1.1467
| Median |  | 1.465 | 1.335 | 1.3736 | 0.9 | 1.3 | 1.3 |  | 1.6 | 1.335 | 1.38 | 1.255 | 1.35 |  | 1.2 | 1.2 | 1.245 | 5.385 | 1.5 | 1.0917 | 1.28 | 1.315 | 1.285 | 1.34 | 1.45 | 1.145
| Std Dev. |  | 0.034448 | 0.030984 | 0.10688 | 0.43012 | 0.08165 | 0.030605 |  | 0.13663 | 0.024221 | 0.074207 | 0.080353 | 0.054924 |  | 0.040825 | 0.040825 | 0.030605 | 1.184 | 0.2582 | 0.034501 | 0.042895 | 0.031345 | 0.092448 | 0.026394 | 0.054772 | 0.056804
| Rel.Std.Dev. |  | 2.37% | 2.33% | 7.611% | 61.45% | 6.446% | 2.351% |  | 8.721% | 1.817% | 5.483% | 6.286% | 4.104% |  | 3.45% | 3.45% | 2.471% | 20.99% | 19.36% | 3.148% | 3.378% | 2.382% | 7.021% | 1.958% | 3.777% | 4.954%
| PDM3 |  | 11.23% | 1.788% | 7.473% | -46.43% | -3.059% | -0.3808% |  | 19.9% | 2.043% | 3.573% | -2.167% | 2.425% |  | -9.437% | -9.437% | -5.228% | 331.6% | 2.043% | -16.14% | -2.804% | 0.7034% | 0.7672% | 3.191% | 10.97% | -12.24%
| Z-Score (Absolute) |  | 1.39 | 0 | 0.83 | NA | 0.71 | 0.32 | 13.15 | NA | 0.04 | 0.26 | 0.58 | 0.09 | 13.15 | NA | NA | 1.03 | 48.44 | NA | 2.63 | 0.67 | 0.16 | 0.15 | 0.21 | NA | 2.06
Results from laboratories 1.07, 1.13, 1.22, 1.23, 1.27 and 1.33 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 233b (Certified Value 37.4 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 40.3 | 38 | 29.77 | 37.0272 | 36.6744 | 37 | 44.8 | 37 | 37.9 | 37 | 33 | 39.8 | 35.5 | 37.3 | 34.5 | 35 | 37 | 35 | 37.2 | 40 | 38.5723 | 39 | 34.4693 | 38.4 | 39.4 | 38.8 | 36.6 |  |  | 1
| 1 | 2 | 39.9 | 37.9 | 32.9 | 37.2 | 37.1736 | 38 | 42.6 | 36 | 37.9 | 36 | 34.3 | 42 | 34.6 | 40.8 | 36.2 | 36 | 34 | 36 | 37.2 | 41 | 39.5854 | 39.6 | 35.1671 | 38.4 | 38.5 | 37.9 | 34.4 |  |  | 33
| 1 | 3 | 40.1 | 38 | 32.14 | 37.2192 | 36.7928 | 37 | 42.8 | 37 | 37.8 | 36 | 31.8 | 41.4 | 39.5 | 39.9 | 36.5 | 35 | 36 | 36 | 38.2 | 39 | 40.2189 | 39.1 | 34.7075 | 38.7 | 38.9 | 39.5 | 33.1 |  |  | 16
| 1 | 4 | 39.3 | 38.3 | 34.61 | 37.3632 | 37.1617 | 38 | 41.8 | 36 | 38 | 37 | 35.2 | 40.8 | 36.2 | 37.2 | 37.1 | 36 | 35 | 36 | 37.6 | 42 | 38.2152 | 38.7 | 35.147 | 38.3 | 38.8 | 40.7 | 32.5 |  |  | 37.3814
| 1 | 5 | 39.4 | 38.4 | 33.35 | 37.5072 | 37.8952 | 37 | 41.3 | 36 | 37.1 | 36 | 34.5 | 40.6 | 36.4 | 38.3 | 36.6 | 37 | 35 | 36 | 37 | 42 | 37.5062 | 39.5 | 34.4296 | 38.7 | 37.9 | 40.9 | 33.8 |  |  | 56
| 1 | 6 | 40.5 | 37.9 | 31.44 | 37.0752 | 35.3297 | 37 | 40.9 | 37 | 37.9 | 36 | 34.9 | 40.8 | 37.6 | 36.2 | 36.8 | 36 | 34 | 36 | 37 | 40 | 38.1619 | 39.1 | 34.2385 | 39.1 | 38.9 | 40 | 31.9
| Mean |  | 39.9167 | 38.0833 | 32.3683 | 37.232 | 36.8379 | 37.3333 | 42.3667 | 36.5 | 37.7667 | 36.3333 | 33.95 | 40.9 | 36.6333 | 38.2833 | 36.2833 | 35.8333 | 35.1667 | 35.8333 | 37.3667 | 40.6667 | 38.71 | 39.1667 | 34.6932 | 38.6 | 38.7333 | 39.6333 | 33.7167
| Median |  | 40 | 38 | 32.52 | 37.2096 | 36.9772 | 37 | 42.2 | 36.5 | 37.9 | 36 | 34.4 | 40.8 | 36.3 | 37.8 | 36.55 | 36 | 35 | 36 | 37.2 | 40.5 | 38.3937 | 39.1 | 34.5884 | 38.55 | 38.85 | 39.75 | 33.45
| Std Dev. |  | 0.48339 | 0.2137 | 1.6695 | 0.17926 | 0.85311 | 0.5164 | 1.3981 | 0.54772 | 0.33267 | 0.5164 | 1.2973 | 0.74565 | 1.7212 | 1.7566 | 0.92394 | 0.75277 | 1.169 | 0.40825 | 0.46332 | 1.2111 | 1.0056 | 0.33267 | 0.38913 | 0.29665 | 0.50067 | 1.1483 | 1.6702
| Rel.Std.Dev. |  | 1.211% | 0.5611% | 5.158% | 0.4815% | 2.316% | 1.383% | 3.3% | 1.501% | 0.8808% | 1.421% | 3.821% | 1.823% | 4.699% | 4.588% | 2.546% | 2.101% | 3.324% | 1.139% | 1.24% | 2.978% | 2.598% | 0.8494% | 1.122% | 0.7685% | 1.293% | 2.897% | 4.954%
| PDM3 |  | 6.782% | 1.878% | -13.41% | -0.3997% | -1.454% | -0.1286% | 13.34% | -2.358% | 1.031% | -2.804% | -9.179% | 9.413% | -2.001% | 2.413% | -2.937% | -4.141% | -5.925% | -4.141% | -0.0394% | 8.789% | 3.554% | 4.776% | -7.191% | 3.26% | 3.617% | 6.024% | -9.804%
| Z-Score (Absolute) |  | 1.24 | 0.36 | 2.39 | 0.05 | 0.24 | 0 | 2.42 | 0.4 | 0.21 | 0.48 | 1.63 | 1.72 | 0.34 | 0.46 | 0.51 | 0.72 | 1.04 | 0.72 | 0.02 | 1.61 | 0.66 | 0.88 | 1.27 | 0.61 | 0.67 | 1.11 | 1.74
| Analytical results for Se in OREAS 233b (Indicative Value 0.92 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 2 | 0.7 | < 1.5 | 0.5 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.07 | < 5 | 1 | 0.5951 | 2.3 | < 2 | < 1 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 2 | 0.6 | < 1.5 | 0.5 | < 1 | < 0.5 | < 5 | 1 | < 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.28 | < 5 | 1 | 0.5745 | 1.1 | < 2 | < 1 | < 0.5 | [6 empty] | 7
| 1 | 3 | < 2 | 0.7 | < 1.5 | 0.6 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 8.23 | < 5 | 1 | 0.5715 | 2.5 | < 2 | < 1 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 2 | 0.6 | < 1.5 | 0.4 | < 1 | 0.7 | < 5 | 1 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 4.04 | < 5 | 1 | 0.561 | 3 | < 2 | < 1 | < 0.5 | [6 empty] | 0.92201
| 1 | 5 | < 2 | 0.5 | < 1.5 | 0.8 | < 1 | 0.5 | < 5 | 1 | < 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.5 | < 5 | 1 | 0.5465 | 3.2 | < 2 | < 1 | < 0.5 | [6 empty] | 9
| 1 | 6 | < 2 | 0.8 | < 1.5 | 0.4 | < 1 | 0.7 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.4 | < 5 | 1 | 0.572 | 2.5 | < 2 | < 1 | < 0.5
| Mean |  | 1.02: 0.65; 1.04: 0.53333; 1.07: 0.63333; 1.10: 1; 1.11: 1; 1.12: 1; 1.13: 1.8333; 1.18: 1; 1.19: 6.2533; 1.23: 1; 1.24: 0.5701; 1.27: 2.4333
| Median |  | 1.02: 0.65; 1.04: 0.5; 1.07: 0.7; 1.10: 1; 1.11: 1; 1.12: 1; 1.13: 2; 1.18: 1; 1.19: 6.34; 1.23: 1; 1.24: 0.57175; 1.27: 2.5
| Std Dev. |  | 1.02: 0.10488; 1.04: 0.15055; 1.07: 0.11547; 1.10: 0; 1.11: 0; 1.12: 0; 1.13: 0.40825; 1.18: 0; 1.19: 1.3359; 1.23: 0; 1.24: 0.016061; 1.27: 0.73666
| Rel.Std.Dev. |  | 1.02: 16.14%; 1.04: 28.23%; 1.07: 18.23%; 1.10: 0%; 1.11: 0%; 1.12: 0%; 1.13: 22.27%; 1.18: 0%; 1.19: 21.36%; 1.23: 0%; 1.24: 2.817%; 1.27: 30.27%
| PDM3 |  | 1.02: -29.5%; 1.04: -42.16%; 1.07: -31.31%; 1.10: 8.459%; 1.11: 8.459%; 1.12: 8.459%; 1.13: 98.84%; 1.18: 8.459%; 1.19: 578.2%; 1.23: 8.459%; 1.24: -38.17%; 1.27: 163.9%
| Z-Score (Absolute) |  | 0 | 0.51 | 0.36 | 0.67 | 0.72 | 0.81 | 2.17 | 0 | 0.24 | 0 | 1.2 | 2.17 | 0 | 0.72 | 0 | 7.59 | 2.17 | 0 | 0.62 | 2.07 | 0 | 0.72 | 1.08
| Analytical results for Sm in OREAS 233b (Certified Value 2.47 ppm) | [30 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*OES; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 2.6; 1.02: 1.89; 1.03: 2.3063; 1.04: 2.7; 1.05: 2.6; 1.07: 2.58; 1.08: 2.6; 1.10: 2.5531 |  |  | 1
| 1 | 2 | 1.01: 2.42; 1.02: 2.07; 1.03: 2.285; 1.04: 2.5; 1.05: 2.5; 1.07: 2.53; 1.08: 2.6; 1.10: 2.4294 |  |  | 18
| 1 | 3 | 1.01: 2.54; 1.02: 2.03; 1.03: 2.2683; 1.04: 2.4; 1.05: 2.5; 1.07: 2.61; 1.08: 2.4; 1.10: 2.4445 |  |  | 16
| 1 | 4 | 1.01: 2.47; 1.02: 2.19; 1.03: 2.3618; 1.04: 2.4; 1.05: 2.5; 1.07: 2.66; 1.08: 2.7; 1.10: 2.425 |  |  | 2.4659
| 1 | 5 | 1.01: 2.43; 1.02: 2.07; 1.03: 2.324; 1.04: 2.3; 1.05: 2.5; 1.07: 2.54; 1.08: 2.4; 1.10: 2.4423 |  |  | 57
| 1 | 6 | 1.01: 2.51; 1.02: 1.94; 1.03: 2.3386; 1.04: 2.2; 1.05: 2.4; 1.07: 2.57; 1.08: 2.4; 1.10: 2.4451
| Mean |  | 1.01: 2.495; 1.02: 2.0317; 1.03: 2.314; 1.04: 2.4167; 1.05: 2.5; 1.07: 2.5817; 1.08: 2.5167; 1.10: 2.4566
| Median |  | 1.01: 2.49; 1.02: 2.05; 1.03: 2.3151; 1.04: 2.4; 1.05: 2.5; 1.07: 2.575; 1.08: 2.5; 1.10: 2.4434
| Std Dev. |  | 1.01: 0.06892; 1.02: 0.10629; 1.03: 0.034584; 1.04: 0.17224; 1.05: 0.063246; 1.07: 0.047924; 1.08: 0.13292; 1.10: 0.048049
| Rel.Std.Dev. |  | 1.01: 2.762%; 1.02: 5.231%; 1.03: 1.495%; 1.04: 7.127%; 1.05: 2.53%; 1.07: 1.856%; 1.08: 5.281%; 1.10: 1.956%
| PDM3 |  | 1.01: 1.181%; 1.02: -17.61%; 1.03: -6.16%; 1.04: -1.996%; 1.05: 1.383%; 1.07: 4.695%; 1.08: 2.059%; 1.10: -0.3786%
| Z-Score (Absolute) |  | 1.01: 0.26; 1.02: 5.99; 1.03: 2.18; 1.04: 0.8; 1.05: 0.33; 1.07: 1.43; 1.08: 0.55; 1.10: 0.26
| Analytical results for Sn in OREAS 233b (Certified Value 1.13 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.4 | 1.1 | 5.67 | < 2.5 | 1 | 1.28 | 1.6 | 1.1 | < 10 | 1.1 | 1 | 1.2 | 1 | 1 | < 10 | 0.9 | 0.9 | 1.1 | 1.29 | 1 | 0.94583 | 0.9 | 0.9429 | 1.2 | 1.2 | 1.3 | 1.6 |  |  | 1
| 1 | 2 | 1.4 | 1.1 | 4.47 | < 2.5 | 1 | 1.23 | 1.5 | 1.1 | < 10 | 1.1 | 1.1 | 1.1 | 1 | 1.1 | < 10 | 0.9 | 0.9 | 1.1 | 1.69 | 1 | 0.91629 | 1.1 | 0.9033 | 1.1 | 1.2 | 1.3 | 1.5 |  |  | 19
| 1 | 3 | 1.4 | 1.1 | 2.35 | < 2.5 | 1 | 1.22 | 1.5 | 1.1 | < 10 | 1.2 | 1 | 1 | 1.1 | 1.1 | < 10 | 1 | 0.9 | 1.1 | 1.88 | 1 | 0.89186 | 1.1 | 0.9001 | 1.2 | 1.1 | 1.3 | 1.4 |  |  | 16
| 1 | 4 | 1.3 | 1.1 | 1.66 | < 2.5 | < 1 | 1.2 | 1.5 | 1.1 | < 10 | 1.2 | 1 | 1.2 | 1 | 1 | < 10 | 0.9 | 1 | 1.1 | 0.93 | 2 | 0.93874 | 1.4 | 0.7385 | 1.1 | 1.2 | 1.4 | 1.4 |  |  | 1.1287
| 1 | 5 | 1.4 | 1.1 | 2.82 | < 2.5 | < 1 | 1.22 | 1.6 | 1.1 | < 10 | 1.2 | 1 | 1.2 | 1.1 | 1.1 | < 10 | 1 | 1 | 1 | 3.21 | 1 | 0.92291 | 1 | 0.6976 | 1.6 | 1.1 | 1.3 | 1.5 |  |  | 58
| 1 | 6 | 1.4 | 1.1 | < 0.1 | < 2.5 | < 1 | 1.24 | 1.5 | 1 | < 10 | 1.2 | 1 | 1.1 | 1 | 1.1 | < 10 | 1.1 | 0.9 | 1.1 | 1.58 | 1 | 0.95506 | 1 | 0.7857 | 1.3 | 1.3 | 1.3 | 1.5
| Mean |  | 1.3833 | 1.1 | 3.394 |  | 1 | 1.2317 | 1.5333 | 1.0833 |  | 1.1667 | 1.0167 | 1.1333 | 1.0333 | 1.0667 |  | 0.96667 | 0.93333 | 1.0833 | 1.7633 | 1.1667 | 0.92845 | 1.0833 | 0.82802 | 1.25 | 1.1833 | 1.3167 | 1.4833
| Median |  | 1.4 | 1.1 | 2.82 |  | 1 | 1.225 | 1.5 | 1.1 |  | 1.2 | 1 | 1.15 | 1 | 1.1 |  | 0.95 | 0.9 | 1.1 | 1.635 | 1 | 0.93083 | 1.05 | 0.8429 | 1.2 | 1.2 | 1.3 | 1.5
| Std Dev. |  | 0.040825 | 2.4324e-16 | 1.6404 |  | 0 | 0.027142 | 0.05164 | 0.040825 |  | 0.05164 | 0.040825 | 0.08165 | 0.05164 | 0.05164 |  | 0.08165 | 0.05164 | 0.040825 | 0.78281 | 0.40825 | 0.02295 | 0.17224 | 0.10087 | 0.18708 | 0.075277 | 0.040825 | 0.075277
| Rel.Std.Dev. |  | 2.951% | 2.211e-14% | 48.33% |  | 0% | 2.204% | 3.368% | 3.768% |  | 4.426% | 4.016% | 7.204% | 4.997% | 4.841% |  | 8.447% | 5.533% | 3.768% | 44.39% | 34.99% | 2.472% | 15.9% | 12.18% | 14.97% | 6.361% | 3.101% | 5.075%
| PDM3 |  | 22.56% | -2.543% | 200.7% |  | -11.4% | 9.123% | 35.85% | -4.019% |  | 3.364% | -9.926% | 0.4107% | -8.449% | -5.496% |  | -14.36% | -17.31% | -4.019% | 56.23% | 3.364% | -17.74% | -4.019% | -26.64% | 10.75% | 4.841% | 16.65% | 31.42%
| Z-Score (Absolute) |  | 0.97 | 0.3 | 7.51 | 0.37 | NA | 0.29 | 1.65 | 0.37 | 17.23 | 0 | 0.67 | 0.15 | 0.6 | 0.45 | 17.23 | 0.9 | 1.05 | 0.37 | 2.68 | NA | 1.07 | 0.37 | 1.52 | 0.37 | 0.07 | 0.67 | 1.42
Results from laboratory 1.07 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 233b (Certified Value 198 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 186.2 | 197.64 | 209 | 195.9608 | 197.3113 | 202 | 204 | 204 | 193.05 | 198 | 189 | 209 | 201 | 200 | 206 | 202 | 195 | 185 | 197 | 97.72 | 181.3075 | 192 | 198.4277 | 208 | 228.2338 | 197 | 216.6 | 201.5 | 187 | 1
| 1 | 2 | 182.9 | 192.35 | 203 | 196.8918 | 199.9522 | 191 | 204 | 200 | 193.84 | 199 | 191 | 218 | 195.5 | 206 | 204 | 209 | 192.1 | 185 | 197 | 103.59 | 180.405 | 191 | 203.4551 | 205 | 225.2918 | 199 | 202.6 | 202.6 | 173 | 20
| 1 | 3 | 183.7 | 194.76 | 206 | 198.5774 | 196.9097 | 194 | 204 | 201 | 191.1 | 196 | 185 | 216 | 197 | 197.5 | 208 | 207 | 190.6 | 182 | 193 | 94.26 | 188.176 | 191 | 204.7627 | 207 | 225.7586 | 195 | 208.3 | 201.2 | 169 | 16
| 1 | 4 | 184.7 | 194.06 | 187 | 198.8616 | 199.4931 | 187 | 202 | 204 | 192.11 | 192 | 188 | 215 | 196.5 | 199 | 209 | 206 | 195.1 | 183 | 192 | 104.01 | 185.8485 | 192 | 193.3154 | 205 | 224.8702 | 201 | 203.7 | 202 | 169 | 197.5049
| 1 | 5 | 178.8 | 194.98 | 192 | 199.5084 | 203.3853 | 185 | 203 | 201 | 190.19 | 195 | 190 | 216 | 199.5 | 200 | 205 | 211 | 191.4 | 187 | 207 | 76.02 | 182.9035 | 192 | 188.7112 | 204 | 219.2018 | 199 | 208.5 | 202.5 | 173 | 59
| 1 | 6 | 189.2 | 195.83 | 197 | 196.7644 | 189.8578 | 185 | 204 | 209 | 194.08 | 192 | 184 | 214 | 204 | 213 | 208 | 206 | 194.2 | 187 | 193 | 95.52 | 188.423 | 192 | 196.6661 | 207 | 222.2522 | 199 | 205.7 | 201.1 | 166
| Mean |  | 184.25 | 194.9367 | 199 | 197.7607 | 197.8182 | 190.6667 | 203.5 | 203.1667 | 192.395 | 195.3333 | 187.8333 | 214.6667 | 198.9167 | 202.5833 | 206.6667 | 206.8333 | 193.0667 | 184.8333 | 196.5 | 95.1867 | 184.5106 | 191.6667 | 197.5564 | 206 | 224.2681 | 198.3333 | 207.5667 | 201.8167 | 172.8333
| Median |  | 184.2 | 194.87 | 200 | 197.7346 | 198.4022 | 189 | 204 | 202.5 | 192.58 | 195.5 | 188.5 | 215.5 | 198.25 | 200 | 207 | 206.5 | 193.15 | 185 | 195 | 96.62 | 184.376 | 192 | 197.5469 | 206 | 225.081 | 199 | 207 | 201.75 | 171
| Std Dev. |  | 3.4738 | 1.7663 | 8.5088 | 1.4086 | 4.5344 | 6.5929 | 0.83666 | 3.3116 | 1.5477 | 2.9439 | 2.7869 | 3.0768 | 3.2158 | 5.8687 | 1.9664 | 3.0605 | 1.947 | 2.0412 | 5.5767 | 10.2295 | 3.4708 | 0.5164 | 6.0746 | 1.5492 | 3.1334 | 2.0656 | 5.0222 | 0.64936 | 7.4409
| Rel.Std.Dev. |  | 1.885% | 0.9061% | 4.276% | 0.7123% | 2.292% | 3.458% | 0.4111% | 1.63% | 0.8044% | 1.507% | 1.484% | 1.433% | 1.617% | 2.897% | 0.9515% | 1.48% | 1.008% | 1.104% | 2.838% | 10.75% | 1.881% | 0.2694% | 3.075% | 0.752% | 1.397% | 1.041% | 2.42% | 0.3218% | 4.305%
| PDM3 |  | -6.711% | -1.3% | 0.757% | 0.1295% | 0.1586% | -3.462% | 3.035% | 2.867% | -2.587% | -1.1% | -4.897% | 8.689% | 0.7148% | 2.571% | 4.639% | 4.723% | -2.247% | -6.416% | -0.5088% | -51.81% | -6.579% | -2.956% | 0.02605% | 4.301% | 13.55% | 0.4194% | 5.094% | 2.183% | -12.49%
| Z-Score (Absolute) |  | 1.59 | 0.33 | 0.15 | 0 | 0.01 | 0.83 | 0.67 | 0.64 | 0.63 | 0.29 | 1.17 | 1.99 | 0.14 | 0.57 | 1.05 | 1.07 | 0.55 | 1.52 | 0.15 | 12.05 | 1.56 | 0.72 | 0.02 | 0.97 | 3.11 | 0.07 | 1.15 | 0.48 | 2.93
| Analytical results for Ta in OREAS 233b (Certified Value 0.23 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.38 | 0.24 | 0.4036 | < 0.1 | 0.22 | 0.36 | 0.22 | 12 | 0.4 | 0.23 | 0.24 | 0.21 | 0.25 | 0.2 | 0.2 | 0.22 | < 0.5 | 0.19428 | 0.23 | 0.3585 | 0.5 | 0.48 | 0.38 | 0.21 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.24 | 0.38604 | < 0.1 | 0.2 | 0.37 | 0.22 | 12 | 0.3 | 0.22 | 0.23 | 0.22 | 0.25 | 0.19 | 0.2 | 0.22 | < 0.5 | 0.18515 | 0.24 | 0.3419 | 0.45 | 0.55 | 0.37 | 0.22 | [5 empty] | 21
| 1 | 3 | 0.34 | 0.25 | 0.37808 | 0.1 | 0.21 | 0.35 | 0.22 | 12 | 0.4 | 0.23 | 0.21 | 0.2 | 0.27 | 0.2 | 0.2 | 0.22 | < 0.5 | 0.1835 | 0.23 | 0.461 | 0.47 | 0.45 | 0.38 | 0.2 | [5 empty] | 16
| 1 | 4 | 0.39 | 0.25 | 0.36629 | < 0.1 | 0.2 | 0.37 | 0.23 | 12 | 0.4 | 0.23 | 0.23 | 0.21 | 0.26 | 0.18 | 0.2 | 0.23 | < 0.5 | 0.18432 | 0.23 | 0.3357 | 0.79 | 0.47 | 0.37 | 0.2 | [5 empty] | 0.2349
| 1 | 5 | 0.29 | 0.25 | 0.37836 | < 0.1 | 0.2 | 0.36 | 0.23 | 12 | 0.4 | 0.23 | 0.24 | 0.21 | 0.28 | 0.2 | 0.2 | 0.22 | < 0.5 | 0.17326 | 0.23 | 0.3235 | 0.99 | 0.49 | 0.36 | 0.2 | [5 empty] | 60
| 1 | 6 | 0.28 | 0.23 | 0.42071 | < 0.1 | 0.2 | 0.4 | 0.22 | 12 | 0.4 | 0.22 | 0.24 | 0.21 | 0.29 | 0.2 | 0.2 | 0.22 | < 0.5 | 0.18477 | 0.23 | 0.3496 | 0.51 | 0.5 | 0.33 | 0.2
| Mean |  | 0.33 | 0.24333 | 0.38884 | 0.1 | 0.205 | 0.36833 | 0.22333 | 12 | 0.38333 | 0.22667 | 0.23167 | 0.21 | 0.26667 | 0.195 | 0.2 | 0.22167 |  | 0.18421 | 0.23167 | 0.3617 | 0.61833 | 0.49 | 0.365 | 0.205
| Median |  | 0.32 | 0.245 | 0.3822 | 0.1 | 0.2 | 0.365 | 0.22 | 12 | 0.4 | 0.23 | 0.235 | 0.21 | 0.265 | 0.2 | 0.2 | 0.22 |  | 0.18454 | 0.23 | 0.34575 | 0.505 | 0.485 | 0.37 | 0.2
| Std Dev. |  | 0.047329 | 0.008165 | 0.019876 |  | 0.0083666 | 0.017224 | 0.005164 | 0 | 0.040825 | 0.005164 | 0.01169 | 0.0063246 | 0.01633 | 0.0083666 | 3.0405e-17 | 0.0040825 |  | 0.0066786 | 0.0040825 | 0.050089 | 0.22076 | 0.034059 | 0.018708 | 0.0083666
| Rel.Std.Dev. |  | 14.34% | 3.355% | 5.111% |  | 4.081% | 4.676% | 2.312% | 0% | 10.65% | 2.278% | 5.046% | 3.012% | 6.124% | 4.291% | 1.52e-14% | 1.842% |  | 3.625% | 1.762% | 13.85% | 35.7% | 6.951% | 5.126% | 4.081%
| PDM3 |  | 40.49% | 3.592% | 65.54% | -57.43% | -12.73% | 56.81% | -4.923% | 5009% | 63.19% | -3.504% | -1.375% | -10.6% | 13.53% | -16.98% | -14.86% | -5.632% |  | -21.58% | -1.375% | 53.98% | 163.2% | 108.6% | 55.39% | -12.73%
| Z-Score (Absolute) |  | 1.92 | 0.12 | 3.14 | NA | 0.67 | 2.71 | 0.29 | NA | NA | 0.22 | 0.12 | 0.57 | 0.61 | 0.88 | NA | 0.33 | 0.26 | 1.11 | 0.12 | 2.58 | 7.9 | 5.24 | 2.65 | 0.67
Results from laboratories 1.07, 1.13 and 1.23 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 233b (Certified Value 0.53 ppm) | [30 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab | [5 empty] | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.05: 1.03; 1.07: 1.04; 1.08: 1.07; 1.10: 1.08; 1.11: 1.24; 1.12: 1.27; 1.13: 1.28; 1.14: 1.32 | [5 empty] | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.05: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*OES; 1.11: 4A*MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.01: 0.61; 1.02: 0.54; 1.05: 0.47; 1.07: 0.50347; 1.08: 0.6; 1.10: 0.54; 1.11: 0.55; 1.12: 0.45; 1.13: 3.18; 1.14: 0.57 | [5 empty] | 1
| 1 | 2 | 1.01: 0.61; 1.02: 0.57; 1.05: 0.51; 1.07: 0.51738; 1.08: 0.6; 1.10: 0.52; 1.11: 0.53; 1.12: 0.5; 1.13: 3.0667; 1.14: 0.56 | [5 empty] | 22
| 1 | 3 | 1.01: 0.61; 1.02: 0.55; 1.05: 0.51; 1.07: 0.50764; 1.08: 0.6; 1.10: 0.52; 1.11: 0.55; 1.12: 0.5; 1.13: 3.0909; 1.14: 0.54 | [5 empty] | 16
| 1 | 4 | 1.01: 0.59; 1.02: 0.56; 1.05: 0.53; 1.07: 0.5157; 1.08: 0.6; 1.10: 0.53; 1.11: 0.54; 1.12: 0.5; 1.13: 3.1383; 1.14: 0.55 | [5 empty] | 0.53364
| 1 | 5 | 1.01: 0.59; 1.02: 0.53; 1.05: 0.51; 1.07: 0.5072; 1.08: 0.6; 1.10: 0.55; 1.11: 0.54; 1.12: 0.5; 1.13: 3.129; 1.14: 0.53 | [5 empty] | 61
| 1 | 6 | 1.01: 0.6; 1.02: 0.58; 1.05: 0.48; 1.07: 0.50347; 1.08: 0.6; 1.10: 0.5; 1.11: 0.53; 1.12: 0.45; 1.13: 3.2464; 1.14: 0.56
| Mean |  | 1.01: 0.60167; 1.02: 0.555; 1.05: 0.50167; 1.07: 0.50914; 1.08: 0.6; 1.10: 0.52667; 1.11: 0.54; 1.12: 0.48333; 1.13: 3.1419; 1.14: 0.55167
| Median |  | 1.01: 0.605; 1.02: 0.555; 1.05: 0.51; 1.07: 0.50742; 1.08: 0.6; 1.10: 0.525; 1.11: 0.54; 1.12: 0.5; 1.13: 3.1336; 1.14: 0.555
| Std Dev. |  | 1.01: 0.0098319; 1.02: 0.018708; 1.05: 0.022286; 1.07: 0.0060201; 1.08: 0; 1.10: 0.017512; 1.11: 0.0089443; 1.12: 0.02582; 1.13: 0.0645; 1.14: 0.01472
| Rel.Std.Dev. |  | 1.01: 1.634%; 1.02: 3.371%; 1.05: 4.442%; 1.07: 1.182%; 1.08: 0%; 1.10: 3.325%; 1.11: 1.656%; 1.12: 5.342%; 1.13: 2.053%; 1.14: 2.668%
| PDM3 |  | 1.01: 12.75%; 1.02: 4.002%; 1.05: -5.992%; 1.07: -4.591%; 1.08: 12.43%; 1.10: -1.307%; 1.11: 1.191%; 1.12: -9.428%; 1.13: 488.8%; 1.14: 3.377%
| Z-Score (Absolute) |  | 1.01: 1.35; 1.02: 0.33; 1.05: 0.84; 1.07: 0.67; 1.08: NA; 1.10: 0.29; 1.11: 0; 1.12: 1.24; 1.13: 56.86; 1.14: 0.25
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 233b (Certified Value 0.096 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.09 | < 0.2 | 0.14225 | < 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.11 | 0.08 | 0.08 | 0.14 | < 5 | 0.1 | < 0.5 | 0.08 | < 1 | 0.0909 | 0.1 | 0.1643 | 0.01 | 0.12 | 0.15 | 0.08 | [5 empty] | 1
| 1 | 2 | 0.09 | < 0.2 | 0.14819 | < 0.1 | < 0.1 | < 0.2 | 15 | 0.1 | 0.1 | 0.08 | 0.08 | 0.12 | < 5 | 0.1 | < 0.5 | 0.07 | < 1 | 0.09338 | 0.09 | 0.115 | 0.02 | 0.12 | 0.13 | 0.07 | [5 empty] | 23
| 1 | 3 | 0.11 | < 0.2 | 0.14919 | 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.08 | 0.13 | 0.09 | 0.1 | < 5 | 0.1 | < 0.5 | 0.09 | < 1 | 0.0972 | 0.09 | 0.0945 | 0.03 | 0.13 | 0.12 | 0.08 | [5 empty] | 16
| 1 | 4 | 0.1 | < 0.2 | 0.12474 | < 0.1 | < 0.1 | < 0.2 | 15 | 0.1 | 0.08 | 0.08 | 0.11 | 0.08 | < 5 | 0.1 | < 0.5 | 0.07 | < 1 | 0.09739 | 0.09 | 0.1356 | 0.1 | 0.11 | 0.16 | 0.08 | [5 empty] | 0.095987
| 1 | 5 | 0.08 | < 0.2 | 0.13925 | < 0.1 | 0.2 | < 0.2 | 15 | < 0.1 | 0.12 | 0.09 | 0.09 | 0.17 | < 5 | 0.1 | < 0.5 | 0.08 | < 1 | 0.08578 | 0.1 | 0.1233 | 0.08 | 0.13 | 0.14 | 0.08 | [5 empty] | 62
| 1 | 6 | 0.09 | < 0.2 | 0.13695 | < 0.1 | 0.2 | < 0.2 | 15 | < 0.1 | 0.1 | 0.09 | 0.08 | 0.12 | < 5 | 0.1 | < 0.5 | 0.08 | < 1 | 0.08866 | 0.1 | 0.0781 | < 0.01 | 0.13 | 0.14 | 0.1
| Mean |  | 0.093333 |  | 0.14009 | 0.1 | 0.2 |  | 15 | 0.1 | 0.098333 | 0.091667 | 0.088333 | 0.12167 |  | 0.1 |  | 0.078333 |  | 0.092218 | 0.095 | 0.11847 | 0.048 | 0.12333 | 0.14 | 0.081667
| Median |  | 0.09 |  | 0.14075 | 0.1 | 0.2 |  | 15 | 0.1 | 0.1 | 0.085 | 0.085 | 0.12 |  | 0.1 |  | 0.08 |  | 0.09214 | 0.095 | 0.11915 | 0.03 | 0.125 | 0.14 | 0.08
| Std Dev. |  | 0.010328 |  | 0.008934 |  | 0 |  | 0 | 0 | 0.016021 | 0.019408 | 0.01169 | 0.031252 |  | 1.5202e-17 |  | 0.0075277 |  | 0.004664 | 0.0054772 | 0.030442 | 0.039623 | 0.008165 | 0.014142 | 0.0098319
| Rel.Std.Dev. |  | 11.07% |  | 6.377% |  | 0% |  | 0% | 0% | 16.29% | 21.17% | 13.23% | 25.69% |  | 1.52e-14% |  | 9.61% |  | 5.058% | 5.766% | 25.7% | 82.55% | 6.62% | 10.1% | 12.04%
| PDM3 |  | -2.764% |  | 45.95% | 4.181% | 108.4% |  | 15527% | 4.181% | 2.445% | -4.5% | -7.973% | 26.75% |  | 4.181% |  | -18.39% |  | -3.926% | -1.028% | 23.42% | -49.99% | 28.49% | 45.85% | -14.92%
| Z-Score (Absolute) |  | 0.24 | 0 | 1.46 | NA | NA | 0 | NA | NA | 0.06 | 0.3 | 0.43 | 0.79 | 87.64 | NA | 5.48 | 0.79 | 14.61 | 0.28 | 0.18 | 0.67 | 2.16 | 0.85 | 1.46 | 0.67
Results from laboratories 1.07, 1.10, 1.13 and 1.22 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 233b (Certified Value 0.93 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.4 | 0.95 | 0.88 | 0.8021 | 0.91367 | 1.1 | 1.06 | 1.07 | 0.91 | 0.9 | 0.9 | 0.94 | 0.9 | 0.91 | 0.88 | 0.8 | 1.01 | 1 | 0.9806 | 0.94 | 1.2781 | 1.4 | 1 | 1 | 0.9 |  |  |  |  | 1
| 1 | 2 | 3.4 | 0.94 | 0.87 | 0.81605 | 0.90072 | 1.1 | 1.02 | 1.05 | 0.92 | 0.9 | 0.93 | 0.9 | 0.92 | 0.92 | 0.86 | 0.8 | 1.01 | 1.1 | 0.96872 | 0.95 | 1.2118 | 1.3 | 1 | 1 | 0.8 |  |  |  |  | 24
| 1 | 3 | 2.7 | 0.93 | 0.86 | 0.82177 | 0.8965 | 1 | 1.05 | 1.2 | 0.93 | 0.9 | 0.94 | 0.9 | 0.93 | 0.95 | 0.88 | 0.8 | 1.03 | 1 | 0.94832 | 0.95 | 1.2405 | 1.3 | 1 | 1 | 0.8 |  |  |  |  | 16
| 1 | 4 | 2.4 | 0.96 | 1 | 0.79719 | 0.92919 | 1 | 1.05 | 1.08 | 0.93 | 1 | 0.93 | 0.91 | 0.9 | 0.93 | 0.85 | 0.8 | 1.03 | 1 | 1.0019 | 0.95 | 1.1845 | 1.3 | 1 | 1.1 | 0.8 |  |  |  |  | 0.92626
| 1 | 5 | 2.7 | 0.92 | 0.88 | 0.79819 | 0.95252 | 1 | 1.02 | 1.12 | 0.92 | 1 | 0.93 | 0.93 | 0.88 | 0.92 | 0.89 | 0.8 | 0.99 | 1 | 0.94564 | 0.98 | 1.1668 | 1.4 | 1 | 1.1 | 0.8 |  |  |  |  | 63
| 1 | 6 | 1.9 | 0.98 | 0.83 | 0.8161 | 0.93573 | 1.2 | 1 | 1.15 | 0.92 | 1 | 0.89 | 0.92 | 0.88 | 0.95 | 0.87 | 0.8 | 0.98 | 1 | 0.99728 | 0.93 | 1.1687 | 1.8 | 1 | 1.1 | 0.8
| Mean |  | 2.5833 | 0.94667 | 0.88667 | 0.80857 | 0.92139 | 1.0667 | 1.0333 | 1.1117 | 0.92167 | 0.95 | 0.92 | 0.91667 | 0.90167 | 0.93 | 0.87167 | 0.8 | 1.0083 | 1.0167 | 0.97374 | 0.95 | 1.2084 | 1.4167 | 1 | 1.05 | 0.81667
| Median |  | 2.55 | 0.945 | 0.875 | 0.80908 | 0.92143 | 1.05 | 1.035 | 1.1 | 0.92 | 0.95 | 0.93 | 0.915 | 0.9 | 0.925 | 0.875 | 0.8 | 1.01 | 1 | 0.97466 | 0.95 | 1.1982 | 1.35 | 1 | 1.05 | 0.8
| Std Dev. |  | 0.49565 | 0.021602 | 0.058538 | 0.010637 | 0.021642 | 0.08165 | 0.023381 | 0.056362 | 0.0075277 | 0.054772 | 0.02 | 0.01633 | 0.020412 | 0.016733 | 0.01472 | 1.2162e-16 | 0.020412 | 0.040825 | 0.023887 | 0.016733 | 0.044222 | 0.19408 | 0 | 0.054772 | 0.040825
| Rel.Std.Dev. |  | 19.19% | 2.282% | 6.602% | 1.316% | 2.349% | 7.655% | 2.263% | 5.07% | 0.8168% | 5.766% | 2.174% | 1.781% | 2.264% | 1.799% | 1.689% | 1.52e-14% | 2.024% | 4.016% | 2.453% | 1.761% | 3.66% | 13.7% | 0% | 5.216% | 4.999%
| PDM3 |  | 178.9% | 2.203% | -4.275% | -12.71% | -0.5262% | 15.16% | 11.56% | 20.02% | -0.4963% | 2.563% | -0.6763% | -1.036% | -2.656% | 0.4034% | -5.894% | -13.63% | 8.86% | 9.76% | 5.125% | 2.563% | 30.46% | 52.94% | 7.961% | 13.36% | -11.83%
| Z-Score (Absolute) |  | NA | 0.44 | 0.83 | 2.5 | 0.09 | NA | 2.29 | 3.96 | 0.09 | NA | 0.12 | 0.2 | 0.51 | 0.09 | 1.15 | NA | 1.76 | NA | 1.02 | 0.51 | 6.02 | NA | NA | NA | NA
| Analytical results for Ti in OREAS 233b (Certified Value 0.597 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.61 | 0.6021 | 0.59452 | 0.60131 | 0.502 | 0.6474 | 0.601 | 0.5966 | 0.58 | 0.599 | 0.627 | 0.617 | 0.588 | 0.594 | 0.56 | 0.6115 | 0.588 | 0.602 | 0.631 | 0.54855 | 0.587 | 0.57885 | 0.596 | 0.49629 | 0.58 | 0.61 | 0.6 | 0.7 |  | 1
| 1 | 2 | 0.6 | 0.5976 | 0.59473 | 0.60947 | 0.467 | 0.6414 | 0.588 | 0.6013 | 0.58 | 0.61 | 0.644 | 0.597 | 0.606 | 0.587 | 0.56 | 0.6158 | 0.594 | 0.608 | 0.618 | 0.55003 | 0.585 | 0.60459 | 0.602 | 0.51799 | 0.59 | 0.6 | 0.6 | 0.7 |  | 25
| 1 | 3 | 0.6 | 0.6066 | 0.59114 | 0.60161 | 0.538 | 0.6474 | 0.596 | 0.596 | 0.57 | 0.589 | 0.647 | 0.592 | 0.582 | 0.608 | 0.56 | 0.6081 | 0.596 | 0.597 | 0.628 | 0.57424 | 0.583 | 0.61684 | 0.599 | 0.50789 | 0.58 | 0.61 | 0.6 | 0.7 |  | 16
| 1 | 4 | 0.59 | 0.6001 | 0.59759 | 0.61129 | 0.221 | 0.6414 | 0.598 | 0.6069 | 0.57 | 0.602 | 0.644 | 0.598 | 0.587 | 0.603 | 0.56 | 0.6196 | 0.594 | 0.58 | 0.618 | 0.5626 | 0.586 | 0.58024 | 0.589 | 0.50829 | 0.59 | 0.6 | 0.6 | 0.7 |  | 0.59663
| 1 | 5 | 0.59 | 0.596 | 0.5985 | 0.61918 | 0.265 | 0.6414 | 0.6 | 0.5964 | 0.57 | 0.599 | 0.643 | 0.597 | 0.574 | 0.59 | 0.58 | 0.6093 | 0.594 | 0.634 | 0.63 | 0.55598 | 0.583 | 0.56774 | 0.593 | 0.50239 | 0.59 | 0.58 | 0.6 | 0.7 |  | 64
| 1 | 6 | 0.6 | 0.6059 | 0.59479 | 0.57682 | 0.222 | 0.6474 | 0.612 | 0.605 | 0.55 | 0.591 | 0.631 | 0.576 | 0.622 | 0.596 | 0.56 | 0.6174 | 0.586 | 0.587 | 0.627 | 0.57197 | 0.586 | 0.57663 | 0.599 | 0.51409 | 0.59 | 0.6 | 0.59 | 0.7
| Mean |  | 0.59833 | 0.60138 | 0.59521 | 0.60328 | 0.36917 | 0.6444 | 0.59917 | 0.60037 | 0.57 | 0.59833 | 0.63933 | 0.59617 | 0.59317 | 0.59633 | 0.56333 | 0.61362 | 0.592 | 0.60133 | 0.62533 | 0.56056 | 0.585 | 0.58748 | 0.59633 | 0.50782 | 0.58667 | 0.6 | 0.59833 | 0.7
| Median |  | 0.6 | 0.6011 | 0.59476 | 0.60554 | 0.366 | 0.6444 | 0.599 | 0.59895 | 0.57 | 0.599 | 0.6435 | 0.597 | 0.5875 | 0.595 | 0.56 | 0.61365 | 0.594 | 0.5995 | 0.6275 | 0.55929 | 0.5855 | 0.57955 | 0.5975 | 0.50809 | 0.59 | 0.6 | 0.6 | 0.7
| Std Dev. |  | 0.0075277 | 0.0043144 | 0.0026031 | 0.014574 | 0.14845 | 0.0032863 | 0.0078081 | 0.0047752 | 0.010954 | 0.0076333 | 0.0082138 | 0.013136 | 0.017623 | 0.0079162 | 0.008165 | 0.0046568 | 0.004 | 0.018928 | 0.0058538 | 0.010931 | 0.0016733 | 0.018914 | 0.0047188 | 0.0078151 | 0.005164 | 0.010954 | 0.0040825 | 1.2162e-16
| Rel.Std.Dev. |  | 1.258% | 0.7174% | 0.4373% | 2.416% | 40.21% | 0.51% | 1.303% | 0.7954% | 1.922% | 1.276% | 1.285% | 2.203% | 2.971% | 1.327% | 1.449% | 0.7589% | 0.6757% | 3.148% | 0.9361% | 1.95% | 0.286% | 3.219% | 0.7913% | 1.539% | 0.8802% | 1.826% | 0.6823% | 1.737e-14%
| PDM3 |  | 0.2854% | 0.7966% | -0.2378% | 1.115% | -38.12% | 8.007% | 0.4251% | 0.6262% | -4.463% | 0.2854% | 7.157% | -0.07772% | -0.5805% | -0.04978% | -5.581% | 2.847% | -0.7761% | 0.7883% | 4.811% | -6.045% | -1.949% | -1.533% | -0.04978% | -14.89% | -1.67% | 0.5648% | 0.2854% | 17.33%
| Z-Score (Absolute) |  | 0.12 | 0.48 | 0.25 | 0.71 | 27.27 | 5.63 | 0.22 | 0.36 | 3.27 | 0.12 | 5.02 | 0.14 | 0.5 | 0.12 | 4.06 | 1.95 | 0.64 | 0.48 | 3.35 | 4.4 | 1.47 | 1.18 | 0.12 | 10.7 | 1.28 | 0.32 | 0.12 | 12.27
| Analytical results for Tl in OREAS 233b (Certified Value 0.21 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.19 | 0.21 | 0.21826 | 0.21 | 0.24 | 0.21 | 0.1 | 0.2 | 0.24 | 0.2 | 0.21 | < 5 | 0.21 | < 0.5 | 0.22 | < 0.1 | 0.2 | 0.21676 | 0.22 | 0.2319 | 0.16 | 0.2 | 0.21 | 0.2 | [5 empty] | 1
| 1 | 2 | 0.19 | 0.21 | 0.21568 | 0.21 | 0.22 | 0.22 | 0.1 | 0.2 | 0.21 | 0.22 | 0.21 | < 5 | 0.19 | < 0.5 | 0.2 | < 0.1 | 0.2 | 0.21435 | 0.22 | 0.2281 | 0.16 | 0.21 | 0.21 | 0.2 | [5 empty] | 26
| 1 | 3 | 0.19 | 0.2 | 0.20859 | 0.2 | 0.21 | 0.22 | 0.1 | 0.22 | 0.2 | 0.18 | 0.21 | < 5 | 0.2 | < 0.5 | 0.21 | 1.94 | 0.2 | 0.21635 | 0.22 | 0.2351 | 0.16 | 0.2 | 0.22 | 0.18 | [5 empty] | 16
| 1 | 4 | 0.19 | 0.22 | 0.22081 | 0.2 | 0.23 | 0.22 | < 0.1 | 0.2 | 0.21 | 0.21 | 0.2 | < 5 | 0.19 | < 0.5 | 0.21 | < 0.1 | 0.2 | 0.21687 | 0.2 | 0.2246 | 0.16 | 0.2 | 0.23 | 0.18 | [5 empty] | 0.20919
| 1 | 5 | 0.18 | 0.22 | 0.22888 | 0.2 | 0.2 | 0.21 | < 0.1 | 0.21 | 0.23 | 0.19 | 0.2 | < 5 | 0.21 | < 0.5 | 0.2 | < 0.1 | 0.2 | 0.20674 | 0.22 | 0.2233 | 0.21 | 0.19 | 0.22 | 0.19 | [5 empty] | 65
| 1 | 6 | 0.19 | 0.22 | 0.27047 | 0.2 | 0.23 | 0.22 | < 0.1 | 0.2 | 0.22 | 0.19 | 0.2 | < 5 | 0.2 | < 0.5 | 0.21 | < 0.1 | 0.2 | 0.21378 | 0.21 | 0.2253 | 0.18 | 0.21 | 0.22 | 0.19
| Mean |  | 0.18833 | 0.21333 | 0.22712 | 0.20333 | 0.22167 | 0.21667 | 0.1 | 0.205 | 0.21833 | 0.19833 | 0.205 |  | 0.2 |  | 0.20833 | 1.94 | 0.2 | 0.21414 | 0.215 | 0.22805 | 0.17167 | 0.20167 | 0.21833 | 0.19
| Median |  | 0.19 | 0.215 | 0.21954 | 0.2 | 0.225 | 0.22 | 0.1 | 0.2 | 0.215 | 0.195 | 0.205 |  | 0.2 |  | 0.21 | 1.94 | 0.2 | 0.21535 | 0.22 | 0.2267 | 0.16 | 0.2 | 0.22 | 0.19
| Std Dev. |  | 0.0040825 | 0.008165 | 0.022246 | 0.005164 | 0.01472 | 0.005164 | 1.6997e-17 | 0.0083666 | 0.01472 | 0.01472 | 0.0054772 |  | 0.0089443 |  | 0.0075277 |  | 3.0405e-17 | 0.0038506 | 0.0083666 | 0.0046164 | 0.020412 | 0.0075277 | 0.0075277 | 0.0089443
| Rel.Std.Dev. |  | 2.168% | 3.827% | 9.795% | 2.54% | 6.64% | 2.383% | 1.7e-14% | 4.081% | 6.742% | 7.422% | 2.672% |  | 4.472% |  | 3.613% |  | 1.52e-14% | 1.798% | 3.891% | 2.024% | 11.89% | 3.733% | 3.448% | 4.708%
| PDM3 |  | -9.971% | 1.98% | 8.568% | -2.801% | 5.963% | 3.573% | -52.2% | -2.004% | 4.37% | -5.191% | -2.004% |  | -4.394% |  | -0.4104% | 827.4% | -4.394% | 2.366% | 2.777% | 9.015% | -17.94% | -3.597% | 4.37% | -9.174%
| Z-Score (Absolute) |  | 1.52 | 0.02 | 0.8 | 0.62 | 0.48 | 0.18 | NA | 0.52 | 0.28 | 0.92 | 0.52 | 137.19 | 0.82 | 2.18 | 0.32 | 9.08 | NA | 0.02 | 0.08 | 0.86 | 2.52 | 0.72 | 0.28 | 1.42
Results from laboratory 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 233b (Certified Value 0.32 ppm) | [30 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.24; 1.08: 1.27; 1.10: 1.28 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.10: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 0.34; 1.02: 0.26; 1.04: 0.29695; 1.05: 0.3; 1.07: 0.31; 1.08: 0.35; 1.10: 0.34533 |  | 1
| 1 | 2 | 1.01: 0.33; 1.02: 0.3; 1.04: 0.293; 1.05: 0.3; 1.07: 0.31; 1.08: 0.3; 1.10: 0.34338 |  | 27
| 1 | 3 | 1.01: 0.33; 1.02: 0.3; 1.04: 0.28938; 1.05: 0.3; 1.07: 0.32; 1.08: 0.3; 1.10: 0.33075 |  | 16
| 1 | 4 | 1.01: 0.33; 1.02: 0.3; 1.04: 0.30595; 1.05: 0.3; 1.07: 0.31; 1.08: 0.35; 1.10: 0.34437 |  | 0.31632
| 1 | 5 | 1.01: 0.33; 1.02: 0.3; 1.04: 0.29151; 1.05: 0.3; 1.07: 0.3; 1.08: 0.35; 1.10: 0.33107 |  | 66
| 1 | 6 | 1.01: 0.33; 1.02: 0.28; 1.04: 0.29401; 1.05: 0.3; 1.07: 0.31; 1.08: 0.3; 1.10: 0.34594
| Mean |  | 1.01: 0.33167; 1.02: 0.29; 1.04: 0.29513; 1.05: 0.3; 1.07: 0.31; 1.08: 0.325; 1.10: 0.34014
| Median |  | 1.01: 0.33; 1.02: 0.3; 1.04: 0.29351; 1.05: 0.3; 1.07: 0.31; 1.08: 0.325; 1.10: 0.34387
| Std Dev. |  | 1.01: 0.0040825; 1.02: 0.016733; 1.04: 0.0058712; 1.05: 0; 1.07: 0.0063246; 1.08: 0.027386; 1.10: 0.0072028
| Rel.Std.Dev. |  | 1.01: 1.231%; 1.02: 5.77%; 1.04: 1.989%; 1.05: 0%; 1.07: 2.04%; 1.08: 8.427%; 1.10: 2.118%
| PDM3 |  | 1.01: 4.851%; 1.02: -8.322%; 1.04: -6.699%; 1.05: -5.16%; 1.07: -1.999%; 1.08: 2.743%; 1.10: 7.529%
| Z-Score (Absolute) |  | 1.01: 0.52; 1.02: 1.02; 1.04: 0.83; 1.05: NA; 1.07: 0.28; 1.08: 0.28; 1.10: 0.84
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 233b (Certified Value 0.3 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.4 | 0.29 | 0.28 | 0.22561 | 0.28257 | 0.3 | 0.43 | 0.3 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.26 | 0.3 | 0.31 | 0.3 | 0.28358 | 0.3 | 0.3317 | 0.296 | 0.4 | 0.3 | 0.266 |  |  |  | 1
| 1 | 2 | 0.5 | 0.29 | 0.29 | 0.22557 | 0.32077 | 0.3 | 0.42 | 0.29 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.25 | 0.3 | 0.3 | 0.3 | 0.27941 | 0.3 | 0.3128 | 0.293 | 0.3 | 0.3 | 0.255 |  |  |  | 28
| 1 | 3 | 0.5 | 0.29 | 0.29 | 0.23632 | 0.29587 | 0.3 | 0.37 | 0.28 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.27 | 0.2 | 0.3 | 0.4 | 0.27284 | 0.3 | 0.317 | 0.288 | 0.3 | 0.3 | 0.235 |  |  |  | 16
| 1 | 4 | 0.4 | 0.3 | 0.32 | 0.21531 | 0.29747 | 0.3 | 0.36 | 0.29 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.25 | 0.3 | 0.31 | 0.3 | 0.28951 | 0.3 | 0.3032 | 0.294 | 0.3 | 0.3 | 0.233 |  |  |  | 0.29707
| 1 | 5 | 0.4 | 0.29 | 0.29 | 0.22666 | 0.30208 | 0.3 | 0.36 | 0.29 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.24 | 0.3 | 0.3 | 0.3 | 0.26892 | 0.3 | 0.2956 | 0.293 | 0.3 | 0.3 | 0.245 |  |  |  | 67
| 1 | 6 | 0.4 | 0.3 | 0.28 | 0.2212 | 0.28103 | 0.3 | 0.34 | 0.28 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.25 | 0.3 | 0.31 | 0.3 | 0.2836 | 0.3 | 0.3032 | 0.287 | 0.3 | 0.3 | 0.253
| Mean |  | 0.43333 | 0.29333 | 0.29167 | 0.22511 | 0.29663 | 0.3 | 0.38 | 0.28833 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 |  | 0.25333 | 0.28333 | 0.305 | 0.31667 | 0.27964 | 0.3 | 0.31058 | 0.29183 | 0.31667 | 0.3 | 0.24783
| Median |  | 0.4 | 0.29 | 0.29 | 0.22559 | 0.29667 | 0.3 | 0.365 | 0.29 | 0.29 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 |  | 0.25 | 0.3 | 0.305 | 0.3 | 0.28149 | 0.3 | 0.308 | 0.293 | 0.3 | 0.3 | 0.249
| Std Dev. |  | 0.05164 | 0.005164 | 0.01472 | 0.006927 | 0.014524 | 0 | 0.036332 | 0.0075277 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.010328 | 0.040825 | 0.0054772 | 0.040825 | 0.0076136 | 0 | 0.012849 | 0.0035449 | 0.040825 | 0 | 0.012656
| Rel.Std.Dev. |  | 11.92% | 1.76% | 5.047% | 3.077% | 4.896% | 0% | 9.561% | 2.611% | 0% | 0% | 0% | 0% | 0% | 0% |  | 4.077% | 14.41% | 1.796% | 12.89% | 2.723% | 0% | 4.137% | 1.215% | 12.89% | 0% | 5.107%
| PDM3 |  | 45.87% | -1.257% | -1.818% | -24.22% | -0.1464% | 0.987% | 27.92% | -2.94% | -2.379% | 0.987% | 0.987% | 0.987% | 0.987% | 0.987% |  | -14.72% | -4.623% | 2.67% | 6.597% | -5.866% | 0.987% | 4.55% | -1.762% | 6.597% | 0.987% | -16.57%
| Z-Score (Absolute) |  | 9.81 | 0.49 | 0.61 | 5.51 | 0.25 | 0 | 5.88 | 0.86 | 0.74 | 0 | 0 | 0 | 0 | 0 | 345.74 | 3.43 | 1.23 | 0.37 | 1.23 | 1.5 | 0 | 0.78 | 0.6 | 1.23 | 0 | 3.84
| Analytical results for V in OREAS 233b (Certified Value 262 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 275 | 276 | 225.55 | 258.7367 | 257.1355 | 245 | 328 | 252 | 247 | 249 | 259 | 274 | 266 | 259 | 260 | 262 | 278 | 261 | 265 | 209.11 | 249.3936 | 260 | 268.0872 | 263 | 235.1805 | 266 | 266 | 279 | 243 | 1
| 1 | 2 | 274 | 259 | 221.62 | 259.9262 | 261.9691 | 214 | 330 | 245 | 260 | 250 | 254 | 281 | 255 | 270 | 261 | 268 | 271 | 266 | 267 | 207.29 | 250.9116 | 255 | 273.2219 | 265 | 240.2946 | 265 | 260 | 283 | 231 | 29
| 1 | 3 | 276 | 274 | 237.61 | 256.8803 | 257.8485 | 246 | 327 | 249 | 264 | 249 | 257 | 280 | 259 | 256 | 263 | 265 | 277 | 266 | 270 | 201.37 | 263.2212 | 255 | 275.2651 | 266 | 239.9844 | 267 | 263 | 275 | 224 | 16
| 1 | 4 | 268 | 273 | 233.92 | 252.0022 | 259.7323 | 125 | 325 | 247 | 273 | 252 | 256 | 279 | 261 | 259 | 263 | 265 | 268 | 266 | 260 | 206.18 | 259.8448 | 255 | 265.5958 | 259 | 241.0888 | 265 | 261 | 278 | 223 | 262.0087
| 1 | 5 | 263 | 261 | 232.02 | 252.9129 | 266.2025 | 165 | 327 | 252 | 258 | 248 | 253 | 280 | 261 | 257 | 259 | 274 | 269 | 264 | 284 | 191.52 | 253.8096 | 260 | 259.2869 | 259 | 235.2563 | 266 | 254 | 278 | 227 | 68
| 1 | 6 | 274 | 265 | 231.01 | 255.7657 | 246.6987 | 130 | 324 | 259 | 250 | 247 | 254 | 278 | 253 | 278 | 263 | 265 | 270 | 264 | 270 | 192.43 | 263.5616 | 260 | 266.0301 | 263 | 236.813 | 268 | 260 | 277 | 219
| Mean |  | 271.6667 | 268 | 230.2883 | 256.0373 | 258.2644 | 187.5 | 326.8333 | 250.6667 | 258.6667 | 249.1667 | 255.5 | 278.6667 | 259.1667 | 263.1667 | 261.5 | 266.5 | 272.1667 | 264.5 | 269.3333 | 201.3167 | 256.7904 | 257.5 | 267.9145 | 262.5 | 238.1029 | 266.1667 | 260.6667 | 278.3333 | 227.8333
| Median |  | 274 | 269 | 231.515 | 256.323 | 258.7904 | 189.5 | 327 | 250.5 | 259 | 249 | 255 | 279.5 | 260 | 259 | 262 | 265 | 270.5 | 265 | 268.5 | 203.775 | 256.8272 | 257.5 | 267.0587 | 263 | 238.3987 | 266 | 260.5 | 278 | 225.5
| Std Dev. |  | 5.0859 | 7.2664 | 5.7949 | 3.1381 | 6.5438 | 55.0445 | 2.137 | 4.9261 | 9.4587 | 1.7224 | 2.2583 | 2.5033 | 4.6655 | 8.8412 | 1.7607 | 4.1352 | 4.2622 | 1.9748 | 8.0911 | 7.6814 | 6.2401 | 2.7386 | 5.7549 | 2.9496 | 2.6671 | 1.169 | 3.9833 | 2.6583 | 8.4479
| Rel.Std.Dev. |  | 1.872% | 2.711% | 2.516% | 1.226% | 2.534% | 29.36% | 0.6538% | 1.965% | 3.657% | 0.6913% | 0.8839% | 0.8983% | 1.8% | 3.36% | 0.6733% | 1.552% | 1.566% | 0.7466% | 3.004% | 3.816% | 2.43% | 1.064% | 2.148% | 1.124% | 1.12% | 0.4392% | 1.528% | 0.9551% | 3.708%
| PDM3 |  | 3.686% | 2.287% | -12.11% | -2.279% | -1.429% | -28.44% | 24.74% | -4.329% | -1.276% | -4.901% | -2.484% | 6.358% | -1.085% | 0.4419% | -0.1942% | 1.714% | 3.877% | 0.9508% | 2.796% | -23.16% | -1.992% | -1.721% | 2.254% | 0.1875% | -9.124% | 1.587% | -0.5122% | 6.231% | -13.04%
| Z-Score (Absolute) |  | 1.02 | 0.68 | 2.83 | 0.43 | 0.22 | 6.81 | 6.16 | 0.93 | 0.19 | 1.07 | 0.48 | 1.67 | 0.14 | 0.23 | 0.08 | 0.54 | 1.07 | 0.36 | 0.81 | 5.52 | 0.36 | 0.29 | 0.67 | 0.17 | 2.1 | 0.51 | 0 | 1.64 | 3.05
| Analytical results for W in OREAS 233b (Certified Value 29.9 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 33 | 30.2 | 31.14 | 23.1861 | 30.2646 | 9.5 | 31.7 | 30.1 | 29.2 | 30 | 7.4 | 32.4 | 28 | 27.9 | 29.2 | 24 | 27 | 27.1 | 30 | 35 | 29.4393 | 31.8 | 31.6734 | 31.9 | 31.2 | 33.1 | 28.6 |  |  | 1
| 1 | 2 | 32.4 | 29.9 | 30.02 | 23.4048 | 30.7346 | 2.9 | 30 | 28.1 | 29.3 | 30 | 6.3 | 33.2 | 26.8 | 30.2 | 30.8 | 26 | 25.7 | 26.7 | 30 | 34 | 29.1012 | 32.9 | 32.3027 | 32.3 | 30.4 | 32.8 | 27.9 |  |  | 30
| 1 | 3 | 32.7 | 29.4 | 34.71 | 23.3544 | 29.4631 | 14.8 | 30.7 | 29.1 | 29.8 | 30 | 8.3 | 33.2 | 27.3 | 26.8 | 31.8 | 25 | 28.3 | 27.5 | 30 | 33 | 29.633 | 31.9 | 32.7241 | 32.4 | 29.9 | 33 | 25.8 |  |  | 16
| 1 | 4 | 32.5 | 30.5 | 30.66 | 23.4812 | 31.024 | 0.7 | 30.6 | 30.2 | 30 | 30 | 5.6 | 33.6 | 28.7 | 27.5 | 31.4 | 26 | 25.8 | 27 | 30 | 34 | 30.229 | 31.7 | 32.724 | 32.1 | 30.1 | 34.2 | 25.8 |  |  | 29.8841
| 1 | 5 | 31.7 | 29.7 | 31.51 | 23.5198 | 32.1655 | 0.8 | 30.4 | 28.9 | 29.9 | 30 | 5.3 | 32.7 | 28.6 | 28.6 | 32.4 | 26 | 26.8 | 26.7 | 32 | 34 | 28.6865 | 32 | 31.8642 | 31.3 | 28.7 | 34.3 | 26.9 |  |  | 69
| 1 | 6 | 32.5 | 29.9 | 32.36 | 23.2161 | 30.5612 | 0.2 | 29.7 | 28.9 | 29.8 | 30 | 6.4 | 32.8 | 31.8 | 27.1 | 31.2 | 25 | 27.4 | 27.1 | 30 | 33 | 29.9812 | 31.8 | 31.9157 | 31.7 | 30.1 | 34.3 | 26.4
| Mean |  | 32.4667 | 29.9333 | 31.7333 | 23.3604 | 30.7022 | 4.8167 | 30.5167 | 29.2167 | 29.6667 | 30 | 6.55 | 32.9833 | 28.5333 | 28.0167 | 31.1333 | 25.3333 | 26.8333 | 27.0167 | 30.3333 | 33.8333 | 29.5117 | 32.0167 | 32.2007 | 31.95 | 30.0667 | 33.6167 | 26.9
| Median |  | 32.5 | 29.9 | 31.325 | 23.3796 | 30.6479 | 1.85 | 30.5 | 29 | 29.8 | 30 | 6.35 | 33 | 28.3 | 27.7 | 31.3 | 25.5 | 26.9 | 27.05 | 30 | 34 | 29.5362 | 31.85 | 32.1092 | 32 | 30.1 | 33.65 | 26.65
| Std Dev. |  | 0.43205 | 0.38297 | 1.6579 | 0.13657 | 0.89361 | 5.9945 | 0.69113 | 0.80104 | 0.33267 | 0 | 1.1256 | 0.43089 | 1.7614 | 1.2416 | 1.093 | 0.8165 | 0.98522 | 0.29944 | 0.8165 | 0.75277 | 0.56646 | 0.4446 | 0.45403 | 0.40866 | 0.81158 | 0.71949 | 1.1454
| Rel.Std.Dev. |  | 1.331% | 1.279% | 5.224% | 0.5846% | 2.911% | 124.5% | 2.265% | 2.742% | 1.121% | 0% | 17.18% | 1.306% | 6.173% | 4.432% | 3.511% | 3.223% | 3.672% | 1.108% | 2.692% | 2.225% | 1.919% | 1.389% | 1.41% | 1.279% | 2.699% | 2.14% | 4.258%
| PDM3 |  | 8.642% | 0.1647% | 6.188% | -21.83% | 2.737% | -83.88% | 2.117% | -2.233% | -0.7277% | 0.3878% | -78.08% | 10.37% | -4.52% | -6.249% | 4.18% | -15.23% | -10.21% | -9.595% | 1.503% | 13.22% | -1.246% | 7.136% | 7.752% | 6.913% | 0.6109% | 12.49% | -9.986%
| Z-Score (Absolute) |  | 0.84 | 0.02 | 0.59 | 2.26 | 0.24 | 8.56 | 0.18 | 0.27 | 0.11 | 0 | 7.97 | 1.01 | 0.5 | 0.67 | 0.39 | 1.59 | 1.08 | 1.01 | 0.11 | 1.3 | 0.17 | 0.69 | 0.75 | 0.66 | 0.02 | 1.23 | 1.05
| Analytical results for Y in OREAS 233b (Certified Value 19.7 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 18.6 | 20.02 | 15.98 | 17.8985 | 18.8498 | 19.7 | 20.9 | 20.6 | 18.95 | 19 | 20.3 | 21.4 | 19.6 | 19.8 | 20.1 | 18 | 19 | 18.4 | 19.5 | 20.6 | 20.1749 | 19.7 | 20.014 | 20.2 | 20.1 | 22 | 17.4 |  |  | 1
| 1 | 2 | 18.5 | 19.77 | 17.7 | 18.3333 | 19.3128 | 19.2 | 20 | 20.5 | 18.86 | 19 | 20.2 | 22 | 19 | 21.1 | 20.3 | 19 | 17.5 | 18.3 | 20.3 | 20.7 | 20.1955 | 20.3 | 19.9544 | 20.4 | 19.5 | 21.4 | 16.4 |  |  | 31
| 1 | 3 | 18.4 | 19.69 | 17.41 | 18.167 | 19.1782 | 19.2 | 20.2 | 20.7 | 19.21 | 18 | 19.5 | 21.9 | 20.1 | 18.6 | 21 | 18 | 19.1 | 19 | 20.3 | 20.1 | 20.1373 | 19.9 | 20.1791 | 20.3 | 19.4 | 21.4 | 15.7 |  |  | 16
| 1 | 4 | 18.3 | 19.78 | 18.49 | 18.261 | 19.2775 | 18.6 | 19.7 | 20.5 | 19 | 19 | 20.5 | 21.4 | 19.4 | 20 | 21.3 | 18 | 17.7 | 17.9 | 19.9 | 21.1 | 19.6897 | 19.8 | 20.1853 | 20.8 | 20.2 | 22.4 | 15.4 |  |  | 19.709
| 1 | 5 | 18.2 | 19.51 | 17.83 | 18.2152 | 19.6928 | 18.3 | 19.7 | 20.3 | 18.5 | 19 | 20.5 | 21.2 | 19.8 | 20.1 | 21.3 | 19 | 18.1 | 18.6 | 19.4 | 20.6 | 19.2807 | 20.6 | 19.7852 | 20.7 | 18.7 | 22.3 | 16.1 |  |  | 70
| 1 | 6 | 18.4 | 19.86 | 16.93 | 17.8917 | 18.1795 | 18.6 | 19.1 | 21 | 18.93 | 18 | 20.3 | 21.2 | 20.4 | 19.2 | 22.1 | 18 | 18.3 | 18.4 | 19.5 | 19.8 | 19.8079 | 20.6 | 19.7733 | 20.6 | 19.1 | 22.3 | 15.1
| Mean |  | 18.4 | 19.7717 | 17.39 | 18.1278 | 19.0818 | 18.9333 | 19.9333 | 20.6 | 18.9083 | 18.6667 | 20.2167 | 21.5167 | 19.7167 | 19.8 | 21.0167 | 18.3333 | 18.2833 | 18.4333 | 19.8167 | 20.4833 | 19.881 | 20.15 | 19.9819 | 20.5 | 19.5 | 21.9667 | 16.0167
| Median |  | 18.4 | 19.775 | 17.555 | 18.1911 | 19.2279 | 18.9 | 19.85 | 20.55 | 18.94 | 19 | 20.3 | 21.4 | 19.7 | 19.9 | 21.15 | 18 | 18.2 | 18.4 | 19.7 | 20.6 | 19.9726 | 20.1 | 19.9842 | 20.5 | 19.45 | 22.15 | 15.9
| Std Dev. |  | 0.14142 | 0.17011 | 0.86021 | 0.18841 | 0.51811 | 0.52026 | 0.60222 | 0.23664 | 0.23267 | 0.5164 | 0.37103 | 0.34881 | 0.49967 | 0.85088 | 0.73326 | 0.5164 | 0.65853 | 0.36148 | 0.41191 | 0.46224 | 0.36146 | 0.40373 | 0.18125 | 0.23664 | 0.57619 | 0.45898 | 0.82321
| Rel.Std.Dev. |  | 0.7686% | 0.8604% | 4.947% | 1.039% | 2.715% | 2.748% | 3.021% | 1.149% | 1.231% | 2.766% | 1.835% | 1.621% | 2.534% | 4.297% | 3.489% | 2.817% | 3.602% | 1.961% | 2.079% | 2.257% | 1.818% | 2.004% | 0.9071% | 1.154% | 2.955% | 2.089% | 5.14%
| PDM3 |  | -6.641% | 0.3182% | -11.77% | -8.023% | -3.182% | -3.935% | 1.138% | 4.521% | -4.062% | -5.288% | 2.576% | 9.172% | 0.03912% | 0.4619% | 6.635% | -6.98% | -7.233% | -6.472% | 0.5465% | 3.929% | 0.8731% | 2.238% | 1.385% | 4.014% | -1.06% | 11.46% | -18.73%
| Z-Score (Absolute) |  | 1.31 | 0.03 | 2.26 | 1.57 | 0.67 | 0.81 | 0.13 | 0.75 | 0.84 | NA | 0.39 | 1.61 | 0.08 | 0 | 1.14 | NA | 1.42 | 1.28 | 0.02 | 0.64 | 0.08 | 0.33 | 0.17 | 0.66 | 0.28 | 2.03 | 3.55
Results from laboratories 1.12 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 233b (Certified Value 2.06 ppm) | [30 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.08; 1.08: 1.24; 1.10: 1.27; 1.11: 1.28; 1.12: 1.32 |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.07: 4A*OES; 1.08: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.01: 2.4; 1.02: 2.07; 1.03: 1.6; 1.04: 1.9746; 1.05: 2.2; 1.07: 2.11; 1.08: 2.2; 1.10: 1.9; 1.11: 2.2233; 1.12: 2.1 |  |  | 1
| 1 | 2 | 1.01: 2.4; 1.02: 2.07; 1.03: 1.83; 1.04: 1.9518; 1.05: 2.1; 1.07: 2; 1.08: 2.2; 1.10: 2; 1.11: 2.1593; 1.12: 2 |  |  | 32
| 1 | 3 | 1.01: 2.4; 1.02: 2.03; 1.03: 1.77; 1.04: 1.9448; 1.05: 2.1; 1.07: 2.03; 1.08: 2.3; 1.10: 2; 1.11: 2.1143; 1.12: 2 |  |  | 16
| 1 | 4 | 1.01: 2.3; 1.02: 2.1; 1.03: 1.98; 1.04: 1.9818; 1.05: 2; 1.07: 2.02; 1.08: 2.3; 1.10: 1.8; 1.11: 2.1775; 1.12: 2 |  |  | 2.0555
| 1 | 5 | 1.01: 2.4; 1.02: 2.05; 1.03: 1.78; 1.04: 1.9692; 1.05: 2; 1.07: 1.96; 1.08: 2.2; 1.10: 2; 1.11: 2.1328; 1.12: 1.9 |  |  | 71
| 1 | 6 | 1.01: 2.3; 1.02: 2.04; 1.03: 1.69; 1.04: 1.92; 1.05: 2.1; 1.07: 1.91; 1.08: 2.1; 1.10: 1.8; 1.11: 2.2435; 1.12: 2
| Mean |  | 1.01: 2.3667; 1.02: 2.06; 1.03: 1.775; 1.04: 1.957; 1.05: 2.0833; 1.07: 2.005; 1.08: 2.2167; 1.10: 1.9167; 1.11: 2.1751; 1.12: 2
| Median |  | 1.01: 2.4; 1.02: 2.06; 1.03: 1.775; 1.04: 1.9605; 1.05: 2.1; 1.07: 2.01; 1.08: 2.2; 1.10: 1.95; 1.11: 2.1684; 1.12: 2
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.025298; 1.03: 0.1288; 1.04: 0.022901; 1.05: 0.075277; 1.07: 0.06775; 1.08: 0.075277; 1.10: 0.098319; 1.11: 0.050501; 1.12: 0.063246
| Rel.Std.Dev. |  | 1.01: 2.182%; 1.02: 1.228%; 1.03: 7.256%; 1.04: 1.17%; 1.05: 3.613%; 1.07: 3.379%; 1.08: 3.396%; 1.10: 5.13%; 1.11: 2.322%; 1.12: 3.162%
| PDM3 |  | 1.01: 15.14%; 1.02: 0.2166%; 1.03: -13.65%; 1.04: -4.793%; 1.05: 1.352%; 1.07: -2.459%; 1.08: 7.838%; 1.10: -6.756%; 1.11: 5.817%; 1.12: -2.702%
| Z-Score (Absolute) |  | 1.01: 2.36; 1.02: 0.19; 1.03: 1.82; 1.04: 0.53; 1.05: 0.36; 1.07: 0.19; 1.08: 1.3; 1.10: 0.82; 1.11: 1.01; 1.12: 0.23
| Analytical results for Zn in OREAS 233b (Certified Value 138 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 143 | 147 | 128.79 | 138.79 | 137.1696 | 132 | 158 | 140 | 139 | 131 | 132 | 150 | 147 | 144 | 146 | 132 | 148 | 135 | 135 | 72.88 | 121.1686 | 140 | 134.5758 | 148 | 177.3227 | 128 | 141 | 128 | 131 | 1
| 1 | 2 | 143 | 147 | 126.08 | 139.24 | 138.7948 | 133 | 148 | 137 | 140 | 128 | 135 | 153 | 143 | 150 | 146 | 136 | 142 | 136 | 136 | 65.16 | 121.4693 | 140 | 140.4391 | 147 | 174.097 | 131 | 135 | 130 | 125 | 33
| 1 | 3 | 141 | 146 | 135.28 | 138.03 | 137.7601 | 130 | 148 | 140 | 138 | 127 | 131 | 154 | 143 | 140 | 147 | 135 | 140 | 136 | 130 | 63.42 | 126.8625 | 130 | 140.1397 | 147 | 177.5132 | 126 | 138 | 129 | 120 | 16
| 1 | 4 | 139 | 147 | 133.57 | 141.52 | 140.0201 | 134 | 150 | 143 | 142 | 129 | 135 | 153 | 142 | 143 | 148 | 134 | 144 | 136 | 130 | 66.99 | 123.7779 | 140 | 136.4672 | 146 | 174.2356 | 129 | 137 | 130 | 118 | 137.7769
| 1 | 5 | 138 | 149 | 133.22 | 140.86 | 141.7422 | 133 | 147 | 140 | 138 | 128 | 135 | 155 | 146 | 142 | 147 | 139 | 144 | 136 | 140 | 69.34 | 122.546 | 140 | 131.5073 | 147 | 172.4835 | 128 | 136 | 130 | 122 | 72
| 1 | 6 | 142 | 148 | 134.92 | 136.71 | 132.0568 | 131 | 145 | 143 | 140 | 129 | 131 | 153 | 144 | 151 | 148 | 132 | 146 | 133 | 133 | 65.44 | 126.8625 | 140 | 135.8853 | 147 | 180.2403 | 131 | 139 | 128 | 119
| Mean |  | 141 | 147.3333 | 131.9767 | 139.1917 | 137.9239 | 132.1667 | 149.3333 | 140.5 | 139.5 | 128.6667 | 133.1667 | 153 | 144.1667 | 145 | 147 | 134.6667 | 144 | 135.3333 | 134 | 67.205 | 123.7811 | 138.3333 | 136.5024 | 147 | 175.982 | 128.8333 | 137.6667 | 129.1667 | 122.5
| Median |  | 141.5 | 147 | 133.395 | 139.015 | 138.2775 | 132.5 | 148 | 140 | 139.5 | 128.5 | 133.5 | 153 | 143.5 | 143.5 | 147 | 134.5 | 144 | 136 | 134 | 66.215 | 123.162 | 140 | 136.1763 | 147 | 175.7791 | 128.5 | 137.5 | 129.5 | 121
| Std Dev. |  | 2.0976 | 1.0328 | 3.7037 | 1.7819 | 3.3091 | 1.472 | 4.5461 | 2.2583 | 1.5166 | 1.3663 | 2.0412 | 1.6733 | 1.9408 | 4.4721 | 0.89443 | 2.6583 | 2.8284 | 1.2111 | 3.8471 | 3.4186 | 2.5568 | 4.0825 | 3.3992 | 0.63246 | 2.8677 | 1.9408 | 2.1602 | 0.98319 | 4.8477
| Rel.Std.Dev. |  | 1.488% | 0.701% | 2.806% | 1.28% | 2.399% | 1.114% | 3.044% | 1.607% | 1.087% | 1.062% | 1.533% | 1.094% | 1.346% | 3.084% | 0.6085% | 1.974% | 1.964% | 0.8949% | 2.871% | 5.087% | 2.066% | 2.951% | 2.49% | 0.4302% | 1.63% | 1.506% | 1.569% | 0.7612% | 3.957%
| PDM3 |  | 2.339% | 6.936% | -4.21% | 1.027% | 0.1067% | -4.072% | 8.388% | 1.976% | 1.251% | -6.612% | -3.346% | 11.05% | 4.638% | 5.243% | 6.694% | -2.257% | 4.517% | -1.774% | -2.741% | -51.22% | -10.16% | 0.4039% | -0.925% | 6.694% | 27.73% | -6.491% | -0.07999% | -6.249% | -11.09%
| Z-Score (Absolute) |  | 0.34 | 1.04 | 0.66 | 0.14 | 0 | 0.64 | 1.27 | 0.29 | 0.17 | 1.03 | 0.53 | 1.67 | 0.69 | 0.79 | 1.01 | 0.36 | 0.67 | 0.29 | 0.44 | 7.85 | 1.57 | 0.05 | 0.16 | 1.01 | 4.22 | 1.01 | 0.03 | 0.97 | 1.71
| Analytical results for Zr in OREAS 233b (Certified Value 50 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 61.9 | 48.2 | 73.84 | 55.1348 | 50.7732 | 45 | 58 | 51.8 | 44.6 | 54 | 55.8 | 50.7 | 46.4 | 50.1 | 45.4 | 41 | 44.9 | 57.7 | 45.2 | 46.7176 | 50 | 67.7585 | 54.8 | 33.1186 | 46.7 | 88.4 | 56.3 | 40.4 |  | 1
| 1 | 2 | 63.5 | 47.1 | 63.43 | 55.1445 | 50.8937 | 42 | 58 | 54.9 | 45.6 | 55 | 57.6 | 51 | 46.9 | 53.7 | 46.7 | 42 | 43.2 | 44.8 | 45.7 | 46.7912 | 50 | 69.0058 | 49.1 | 32.6924 | 48.4 | 83.8 | 56.6 | 37.9 |  | 34
| 1 | 3 | 61.7 | 48.2 | 63 | 57.6859 | 51.3811 | 47 | 56 | 48.4 | 47.6 | 53 | 54.9 | 51.6 | 47 | 45.1 | 53.3 | 40 | 45 | 46.4 | 45.7 | 48.898 | 50 | 70.4648 | 51.7 | 33.2304 | 46.5 | 85.2 | 54.7 | 36.9 |  | 16
| 1 | 4 | 59.6 | 51.6 | 68.94 | 55.0766 | 49.9504 | 17 | 56 | 50 | 45.4 | 55 | 56.5 | 49.5 | 61.9 | 51 | 46.7 | 43 | 46.3 | 47.1 | 44.5 | 48.0056 | 50 | 66.3608 | 50.7 | 32.8115 | 47.5 | 82.6 | 58.5 | 35.8 |  | 50.1978
| 1 | 5 | 58.1 | 51 | 66.34 | 56.066 | 52.0366 | 28 | 56 | 50.1 | 44 | 54 | 57.2 | 50.8 | 49 | 50 | 56.3 | 56 | 43.5 | 43.9 | 48.9 | 46.8188 | 50 | 65.2785 | 53.9 | 32.4571 | 48.5 | 82.6 | 59.7 | 37.3 |  | 73
| 1 | 6 | 57.2 | 47 | 70.73 | 55.1251 | 46.6958 | 19 | 58 | 50.5 | 44.9 | 52 | 56.6 | 49.9 | 49.5 | 50 | 49.8 | 42 | 46.6 | 43.7 | 45.9 | 49.0912 | 50 | 67.467 | 51.4 | 32.6655 | 47.6 | 81.8 | 57.9 | 35.7
| Mean |  | 60.3333 | 48.85 | 67.7133 | 55.7055 | 50.2885 | 33 | 57 | 50.95 | 45.35 | 53.8333 | 56.4333 | 50.5833 | 50.1167 | 49.9833 | 49.7 | 44 | 44.9167 | 47.2667 | 45.9833 | 47.7204 | 50 | 67.7226 | 51.9333 | 32.8293 | 47.5333 | 84.0667 | 57.2833 | 37.3333
| Median |  | 60.65 | 48.2 | 67.64 | 55.1397 | 50.8334 | 35 | 57 | 50.3 | 45.15 | 54 | 56.55 | 50.75 | 48 | 50.05 | 48.25 | 42 | 44.95 | 45.6 | 45.7 | 47.4122 | 50 | 67.6127 | 51.55 | 32.752 | 47.55 | 83.2 | 57.25 | 37.1
| Std Dev. |  | 2.4369 | 1.9756 | 4.2565 | 1.0416 | 1.8906 | 13.4015 | 1.0954 | 2.2206 | 1.2422 | 1.169 | 0.97297 | 0.76267 | 5.9051 | 2.7824 | 4.3132 | 5.9666 | 1.3934 | 5.2872 | 1.5158 | 1.0981 | 0 | 1.8472 | 2.0963 | 0.29292 | 0.83106 | 2.432 | 1.7781 | 1.7282
| Rel.Std.Dev. |  | 4.039% | 4.044% | 6.286% | 1.87% | 3.76% | 40.61% | 1.922% | 4.358% | 2.739% | 2.172% | 1.724% | 1.508% | 11.78% | 5.567% | 8.679% | 13.56% | 3.102% | 11.19% | 3.296% | 2.301% | 0% | 2.728% | 4.037% | 0.8922% | 1.748% | 2.893% | 3.104% | 4.629%
| PDM3 |  | 20.19% | -2.685% | 34.89% | 10.97% | 0.1805% | -34.26% | 13.55% | 1.498% | -9.657% | 7.242% | 12.42% | 0.768% | -0.1617% | -0.4273% | -0.9917% | -12.35% | -10.52% | -5.839% | -8.396% | -4.935% | -0.3941% | 34.91% | 3.457% | -34.6% | -5.308% | 67.47% | 14.12% | -25.63%
| Z-Score (Absolute) |  | 1.58 | 0.19 | 2.71 | 0.87 | 0.04 | 2.62 | 1.07 | 0.14 | 0.72 | 0.58 | 0.98 | 0.08 | 0.01 | 0.01 | 0.06 | 0.93 | 0.79 | 0.43 | 0.63 | 0.36 | 0.01 | 2.71 | 0.29 | 2.65 | 0.39 | 5.22 | 1.11 | 1.95
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C513: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C551: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C569: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C606: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C624: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C642: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C660: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C714: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C897: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1119: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1137: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 233b (Certified Value 0.306 ppm) | [32 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.4 | 0.3 | 0.49 | 0.31696 | 0.347 | 0.35 | 0.3 | 0.27 | < 0.5 | 0.3 | 0.33 | 0.31 | 0.3 | < 10 | 0.3 | 0.2 | < 1 | 0.25 | 0.275 | 0.33 | 0.2 | 0.3 | 0.33214 | 0.31 | 0.321 | 0.25 | 0.3 | 0.29 | 0.32 |  |  | 1
| 1 | 2 | 0.39 | 0.3 | 0.44 | 0.32126 | 0.285 | 0.34 | 0.3 | 0.29 | < 0.5 | 0.4 | 0.3 | 0.32 | 0.29 | < 10 | 0.29 | < 0.2 | < 1 | 0.29 | 0.292 | 0.32 | 0.4 | 0.3 | 0.31868 | 0.31 | 0.31 | 0.29 | 0.31 | 0.3 | 0.32 |  |  | 18
| 1 | 3 | 0.41 | 0.3 | 0.41 | 0.38096 | 0.293 | 0.33 | 0.33 | 0.3 | < 0.5 | 0.3 | 0.34 | 0.32 | 0.3 | < 10 | 0.3 | 0.2 | < 1 | 0.29 | 0.275 | 0.32 | < 0.2 | 0.3 | 0.31358 | 0.31 | 0.299 | 0.29 | 0.33 | 0.3 | 0.31 |  |  | 16
| 1 | 4 | 0.39 | 0.3 | 0.42 | 0.32307 | 0.301 | 0.34 | 0.33 | 0.3 | < 0.5 | 0.3 | 0.31 | 0.32 | 0.31 | < 10 | 0.29 | < 0.2 | < 1 | 0.28 | 0.27 | 0.32 | 0.25 | 0.3 | 0.31353 | 0.32 | 0.339 | 0.27 | 0.33 | 0.3 | 0.31 |  |  | 0.30617
| 1 | 5 | 0.41 | 0.3 | 0.37 | 0.30542 | 0.301 | 0.33 | 0.32 | 0.29 | < 0.5 | 0.3 | 0.32 | 0.31 | 0.3 | < 10 | 0.31 | < 0.2 | < 1 | 0.25 | 0.279 | 0.32 | 1.02 | 0.3 | 0.32138 | 0.31 | 0.306 | 0.29 | 0.32 | 0.3 | 0.31 |  |  | 75
| 1 | 6 | 0.41 | 0.3 | 0.44 | 0.31821 | 0.322 | 0.35 | 0.33 | 0.3 | < 0.5 | 0.3 | 0.3 | 0.32 | 0.3 | < 10 | 0.31 | < 0.2 | < 1 | 0.27 | 0.286 | 0.3 | < 0.2 | 0.3 | 0.29696 | 0.32 | 0.27 | 0.3 | 0.3 | 0.3 | 0.32
| Mean |  | 0.40167 | 0.3 | 0.42833 | 0.32764 | 0.30817 | 0.34 | 0.31833 | 0.29167 |  | 0.31667 | 0.31667 | 0.31667 | 0.3 |  | 0.3 | 0.2 |  | 0.27167 | 0.2795 | 0.31833 | 0.4675 | 0.3 | 0.31605 | 0.31333 | 0.3075 | 0.28167 | 0.315 | 0.29833 | 0.315
| Median |  | 0.405 | 0.3 | 0.43 | 0.31973 | 0.301 | 0.34 | 0.325 | 0.295 |  | 0.3 | 0.315 | 0.32 | 0.3 |  | 0.3 | 0.2 |  | 0.275 | 0.277 | 0.32 | 0.325 | 0.3 | 0.31613 | 0.31 | 0.308 | 0.29 | 0.315 | 0.3 | 0.315
| Std Dev. |  | 0.0098319 | 0 | 0.039707 | 0.02684 | 0.022666 | 0.0089443 | 0.01472 | 0.01169 |  | 0.040825 | 0.01633 | 0.005164 | 0.0063246 |  | 0.0089443 | 0 |  | 0.018348 | 0.0081179 | 0.0098319 | 0.37801 | 0 | 0.011585 | 0.005164 | 0.023072 | 0.018348 | 0.013784 | 0.0040825 | 0.0054772
| Rel.Std.Dev. |  | 2.448% | 0% | 9.27% | 8.192% | 7.355% | 2.631% | 4.624% | 4.008% |  | 12.89% | 5.157% | 1.631% | 2.108% |  | 2.981% | 0% |  | 6.754% | 2.904% | 3.089% | 80.86% | 0% | 3.666% | 1.648% | 7.503% | 6.514% | 4.376% | 1.368% | 1.739%
| PDM3 |  | 31.19% | -2.016% | 39.9% | 7.014% | 0.6517% | 11.05% | 3.972% | -4.737% |  | 3.428% | 3.428% | 3.428% | -2.016% |  | -2.016% | -34.68% |  | -11.27% | -8.711% | 3.972% | 52.69% | -2.016% | 3.225% | 2.339% | 0.434% | -8.004% | 2.884% | -2.56% | 2.884%
| Z-Score (Absolute) |  | 3.9 | 0.67 | 5.09 | 0.57 | 0.31 | 1.12 | 0.15 | 1.05 | 2.92 | NA | 0.07 | 0.07 | 0.67 | 210.61 | 0.67 | NA | 8.32 | 1.95 | 1.6 | 0.15 | 1.35 | NA | 0.05 | 0.07 | 0.34 | 1.5 | 0 | 0.75 | 0
Results from laboratories 1.13, 1.19 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 233b (Certified Value 3.19 wt.%) | [32 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.12 | 2.968 | 2.9 | 3.6354 | 3.1131 | 3.08 | 3.572 | 3.49 | 3.07 | 2.98 | 3.38 | 3.23 | 3.23 | 3.21 | 3.19 | 3.14 | 3.344 | 3.38 | 3.06 | 3.51 | 3.7017 | 2.7271 | 2.7405 | 3.06 | 2.9 | 3.43 | 3.02 | 2.87 |  |  |  | 1
| 1 | 2 | 3.2 | 2.929 | 2.8436 | 3.6573 | 3.1077 | 3.06 | 3.556 | 3.49 | 3.0469 | 2.98 | 3.35 | 3.23 | 3.31 | 3.22 | 3.22 | 3.36 | 3.5084 | 3.393 | 3.22 | 3.46 | 4.0227 | 2.7428 | 2.7575 | 3.04 | 3.05 | 3.39 | 3.11 | 2.91 |  |  |  | 0
| 1 | 3 | 3.24 | 3.028 | 2.7923 | 3.6312 | 3.0521 | 3.16 | 3.588 | 3.55 | 3.0545 | 2.98 | 3.34 | 3.22 | 3.28 | 3.24 | 3.15 | 3.28 | 3.576 | 3.388 | 3.11 | 3.44 | 3.0332 | 2.9109 | 2.696 | 3.04 | 2.94 | 3.37 | 3.06 | 2.91 |  |  |  | 16
| 1 | 4 | 3.23 | 2.956 | 2.7642 | 3.6572 | 3.1642 | 2.94 | 3.408 | 3.51 | 2.8961 | 2.97 | 3.33 | 3.24 | 3.26 | 3.25 | 3.24 | 3.18 | 3.5032 | 3.342 | 2.98 | 3.43 | 3.923 | 2.9744 | 2.718 | 3.03 | 2.96 | 3.38 | 3.02 | 2.84 |  |  |  | 3.188
| 1 | 5 | 3.24 | 2.949 | 2.7779 | 3.6827 | 3.0744 | 2.94 | 3.467 | 3.6 | 3.0134 | 2.98 | 3.27 | 3.25 | 3.27 | 3.25 | 3.33 | 3.14 | 3.4652 | 3.348 | 3.1 | 3.38 | 3.2603 | 2.9214 | 2.645 | 3.02 | 2.9 | 3.46 | 3.06 | 2.86 |  |  |  | 76
| 1 | 6 | 3.21 | 2.936 | 2.8827 | 3.6558 | 3.3362 | 3.19 | 3.604 | 3.61 | 3.0781 | 3.02 | 3.26 | 3.22 | 3.26 | 3.3 | 3.33 | 3.07 | 3.5015 | 3.385 | 3.08 | 3.38 | 3.8557 | 2.9083 | 2.6325 | 3.04 | 2.83 | 3.4 | 3.04 | 2.81
| Mean |  | 3.2067 | 2.961 | 2.8268 | 3.6533 | 3.1413 | 3.0617 | 3.5325 | 3.5417 | 3.0265 | 2.985 | 3.3217 | 3.2317 | 3.2683 | 3.245 | 3.2433 | 3.195 | 3.4831 | 3.3727 | 3.0917 | 3.4333 | 3.6328 | 2.8642 | 2.6982 | 3.0383 | 2.93 | 3.405 | 3.0517 | 2.8667
| Median |  | 3.22 | 2.9525 | 2.818 | 3.6565 | 3.1104 | 3.07 | 3.564 | 3.53 | 3.0507 | 2.98 | 3.335 | 3.23 | 3.265 | 3.245 | 3.23 | 3.16 | 3.5023 | 3.3825 | 3.09 | 3.435 | 3.7787 | 2.9096 | 2.707 | 3.04 | 2.92 | 3.395 | 3.05 | 2.865
| Std Dev. |  | 0.045461 | 0.035654 | 0.057008 | 0.018497 | 0.10284 | 0.10591 | 0.077583 | 0.053821 | 0.067721 | 0.017607 | 0.047081 | 0.01169 | 0.026394 | 0.031464 | 0.073666 | 0.10616 | 0.077039 | 0.021924 | 0.078081 | 0.049666 | 0.39722 | 0.10303 | 0.050694 | 0.013292 | 0.073756 | 0.033912 | 0.033714 | 0.039328
| Rel.Std.Dev. |  | 1.418% | 1.204% | 2.017% | 0.5063% | 3.274% | 3.459% | 2.196% | 1.52% | 2.238% | 0.5898% | 1.417% | 0.3617% | 0.8076% | 0.9696% | 2.271% | 3.323% | 2.212% | 0.6501% | 2.526% | 1.447% | 10.93% | 3.597% | 1.879% | 0.4375% | 2.517% | 0.9959% | 1.105% | 1.372%
| PDM3 |  | 0.5863% | -7.12% | -11.33% | 14.6% | -1.464% | -3.962% | 10.81% | 11.09% | -5.065% | -6.367% | 4.194% | 1.37% | 2.521% | 1.789% | 1.736% | 0.2203% | 9.256% | 5.793% | -3.021% | 7.696% | 13.95% | -10.16% | -15.36% | -4.694% | -8.092% | 6.808% | -4.276% | -10.08%
| Z-Score (Absolute) |  | 0.02 | 0.85 | 1.33 | 1.61 | 0.21 | 0.5 | 1.18 | 1.21 | 0.62 | 0.77 | 0.43 | 0.11 | 0.24 | 0.16 | 0.15 | 0.02 | 1.01 | 0.61 | 0.39 | 0.83 | 1.54 | 1.2 | 1.79 | 0.58 | 0.96 | 0.73 | 0.53 | 1.19
| Analytical results for As in OREAS 233b (Certified Value 59 ppm) | [32 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 1
| 1 | 1 | 61 | 58.48 | 52.64 | 57.54 | 58.6289 | 60.6 | 62.9 | 61 | 59 | 54 | 58.1 | 60.8 | 59.2 | 58.5 | < 100 | 61.3 | 54 | 46.8086 | 55 | 55.1 | 59 | 45.55 | 58.006 | 55 | 63.4368 | 60.8 | 50.797 | 58.1 | 59 | 60.2 | 58.8 | 1
| 1 | 2 | 62 | 57.86 | 51.2 | 58.95 | 58.4049 | 59 | 64.4 | 59 | 59 | 54 | 59 | 60.6 | 60.4 | 60.2 | < 100 | 60.5 | 58 | 48.1282 | 57 | 57.4 | 59 | 44.78 | 58.4568 | 60 | 63.4782 | 62.1 | 55.92 | 60.8 | 61 | 61 | 58.9 | 19
| 1 | 3 | 62 | 58.43 | 52.47 | 59.37 | 63.7169 | 58.4 | 61.4 | 61 | 59 | 53 | 56.6 | 61.8 | 60.5 | 60.3 | < 100 | 59.7 | 56 | 45.551 | 57 | 56.3 | 58 | 44.78 | 60.1956 | 55 | 59.4241 | 61.5 | 56.092 | 59 | 61 | 60.5 | 58 | 16
| 1 | 4 | 61 | 59.2 | 50.5 | 58.48 | 58.9797 | 57.3 | 64.5 | 60 | 57 | 54 | 57.3 | 60.7 | 58.4 | 60.6 | < 100 | 60.6 | 55 | 46.9939 | 55 | 55.3 | 58 | 45.63 | 61.6584 | 55 | 65.1503 | 61.5 | 53.444 | 62.9 | 62 | 59.9 | 58.4 | 58.9956
| 1 | 5 | 62 | 59.51 | 50.79 | 58.72 | 58.1457 | 58.2 | 59.9 | 61 | 58 | 55 | 57 | 61.9 | 58.9 | 60.2 | < 100 | 64.2 | 55 | 45.943 | 55 | 54.9 | 59 | 45.34 | 61.4928 | 60 | 60.5262 | 62.5 | 55.496 | 61.6 | 60 | 59.6 | 58.5 | 77
| 1 | 6 | 62 | 58.28 | 51.6 | 58.25 | 59.0386 | 61.4 | 63.9 | 60 | 59 | 55 | 57.6 | 60.5 | 59.8 | 61.9 | < 100 | 62.7 | 56 | 48.9437 | 55 | 56 | 58 | 44.39 | 62.1736 | 55 | 60.4899 | 61.3 | 53.951 | 60 | 58 | 59.8 | 57.6
| Mean |  | 61.6667 | 58.6267 | 51.5333 | 58.5517 | 59.4858 | 59.15 | 62.8333 | 60.3333 | 58.5 | 54.1667 | 57.6 | 61.05 | 59.5333 | 60.2833 |  | 61.5 | 55.6667 | 47.0614 | 55.6667 | 55.8333 | 58.5 | 45.0783 | 60.3305 | 56.6667 | 62.0843 | 61.6167 | 54.2833 | 60.4 | 60.1667 | 60.1667 | 58.3667
| Median |  | 62 | 58.455 | 51.4 | 58.6 | 58.8043 | 58.7 | 63.4 | 60.5 | 59 | 54 | 57.45 | 60.75 | 59.5 | 60.25 |  | 60.95 | 55.5 | 46.9013 | 55 | 55.65 | 58.5 | 45.06 | 60.8442 | 55 | 61.9815 | 61.5 | 54.7235 | 60.4 | 60.5 | 60.05 | 58.45
| Std Dev. |  | 0.5164 | 0.61272 | 0.87612 | 0.62869 | 2.1003 | 1.554 | 1.8457 | 0.8165 | 0.83666 | 0.75277 | 0.85557 | 0.62849 | 0.84301 | 1.087 |  | 1.6625 | 1.3663 | 1.287 | 1.0328 | 0.93737 | 0.54772 | 0.49942 | 1.7575 | 2.582 | 2.2458 | 0.60139 | 2.0194 | 1.747 | 1.472 | 0.5164 | 0.49261
| Rel.Std.Dev. |  | 0.8374% | 1.045% | 1.7% | 1.074% | 3.531% | 2.627% | 2.937% | 1.353% | 1.43% | 1.39% | 1.485% | 1.029% | 1.416% | 1.803% |  | 2.703% | 2.454% | 2.735% | 1.855% | 1.679% | 0.9363% | 1.108% | 2.913% | 4.556% | 3.617% | 0.976% | 3.72% | 2.892% | 2.446% | 0.8583% | 0.844%
| PDM3 |  | 4.528% | -0.6253% | -12.65% | -0.7524% | 0.831% | 0.2618% | 6.505% | 2.268% | -0.84% | -8.185% | -2.366% | 3.482% | 0.9116% | 2.183% |  | 4.245% | -5.643% | -20.23% | -5.643% | -5.36% | -0.84% | -23.59% | 2.263% | -3.948% | 5.235% | 4.443% | -7.987% | 2.381% | 1.985% | 1.985% | -1.066%
| Z-Score (Absolute) |  | 1.05 | 0 | 2.44 | 0.03 | 0.3 | 0.18 | 1.45 | 0.59 | 0.04 | 1.53 | 0.35 | 0.83 | 0.31 | 0.57 | 2.97 | 0.99 | 1.02 | 3.98 | 1.02 | 0.96 | 0.04 | 4.66 | 0.59 | 0.67 | 1.19 | 1.03 | 1.49 | 0.61 | 0.53 | 0.53 | 0.09
| Analytical results for B in OREAS 233b (Certified Value 23.7 ppm) | [32 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.11 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 19 | 20 | 35.32 | 27.4973 | 35 | 31 | 20 | 20 | 20 | 20 | 21 | 17 | 65.46 | 20 | 10 | 12 | < 10 | 28 | 15 | [12 empty] | 1
| 1 | 2 | 21 | 20 | 36.07 | 26.2073 | 34 | 27 | 20 | 20 | 20 | 20 | 20 | 19 | 61.49 | 20 | 10 | 16 | < 10 | 28 | 15.9 | [12 empty] | 3
| 1 | 3 | 21 | 20 | 36.87 | 27.3458 | 34 | 28 | 20 | 20 | 20 | 20 | 20 | 18 | 67.11 | 10 | 10 | 14 | < 10 | 30 | 15.1 | [12 empty] | 16
| 1 | 4 | 22 | 20 | 35.79 | 24.6521 | 33 | 27 | 20 | 20 | 20 | 20 | 21 | 17 | 68.57 | 10 | 10 | 13 | < 10 | 28 | 14.9 | [12 empty] | 23.6901
| 1 | 5 | 22 | 21 | 34.9 | 24.6809 | 34 | 27 | 20 | 20 | 20 | 20 | 22 | 17 | 67.75 | 10 | 10 | 14 | < 10 | 28 | 15.1 | [12 empty] | 78
| 1 | 6 | 21 | 20 | 36.37 | 24.5463 | 36 | 29 | 20 | 20 | 20 | 20 | 22 | 17 | 87.33 | 10 | 10 | 11 | < 10 | 28 | 14.3
| Mean |  | 1.01: 21; 1.02: 20.1667; 1.03: 35.8867; 1.04: 25.8216; 1.05: 34.3333; 1.07: 28.1667; 1.08: 20; 1.10: 20; 1.11: 20; 1.12: 20; 1.13: 21; 1.14: 17.5; 1.15: 69.6183; 1.16: 13.3333; 1.17: 10; 1.18: 13.3333; 1.21: 28.3333; 1.22: 15.05
| Median |  | 1.01: 21; 1.02: 20; 1.03: 35.93; 1.04: 25.4441; 1.05: 34; 1.07: 27.5; 1.08: 20; 1.10: 20; 1.11: 20; 1.12: 20; 1.13: 21; 1.14: 17; 1.15: 67.43; 1.16: 10; 1.17: 10; 1.18: 13.5; 1.21: 28; 1.22: 15.05
| Std Dev. |  | 1.01: 1.0954; 1.02: 0.40825; 1.03: 0.71248; 1.04: 1.3839; 1.05: 1.0328; 1.07: 1.6021; 1.08: 0; 1.10: 0; 1.11: 0; 1.12: 0; 1.13: 0.89443; 1.14: 0.83666; 1.15: 9.0327; 1.16: 5.164; 1.17: 0; 1.18: 1.7512; 1.21: 0.8165; 1.22: 0.51284
| Rel.Std.Dev. |  | 1.01: 5.216%; 1.02: 2.024%; 1.03: 1.985%; 1.04: 5.359%; 1.05: 3.008%; 1.07: 5.688%; 1.08: 0%; 1.10: 0%; 1.11: 0%; 1.12: 0%; 1.13: 4.259%; 1.14: 4.781%; 1.15: 12.97%; 1.16: 38.73%; 1.17: 0%; 1.18: 13.13%; 1.21: 2.882%; 1.22: 3.408%
| PDM3 |  | 1.01: -11.36%; 1.02: -14.87%; 1.03: 51.48%; 1.04: 8.997%; 1.05: 44.93%; 1.07: 18.9%; 1.08: -15.58%; 1.10: -15.58%; 1.11: -15.58%; 1.12: -15.58%; 1.13: -11.36%; 1.14: -26.13%; 1.15: 193.9%; 1.16: -43.72%; 1.17: -57.79%; 1.18: -43.72%; 1.21: 19.6%; 1.22: -36.47%
| Z-Score (Absolute) |  | 0 | 0.08 | 1.4 | 0.45 | 1.25 | 0.67 | NA | NA | NA | NA | 0 | 0.33 | 4.57 | NA | NA | 0.72 | 1.51 | 0.69 | 0.56
| Analytical results for Ba in OREAS 233b (Certified Value 39.1 ppm) | [32 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 38 | 39.07 | 42.36 | 46.86 | 41.1297 | 40.7 | 49 | 46 | 39 | 40 | 39 | 40 | 40 | 40 | 52.64 | 40 | 26 | 39.2694 | 38 | 31.9 | 39 | 35.27 | 36.225 | 32 | 40 | 38 | 41 | 35 | 36 |  |  | 1
| 1 | 2 | 38 | 38.24 | 40.63 | 46.2 | 40.5673 | 39.5 | 48 | 45 | 38 | 39 | 40 | 40 | 40 | 40 | 53.561 | 40 | 28 | 41.2876 | 40 | 32.4 | 40 | 33.78 | 36.729 | 32 | 40 | 39 | 43 | 36 | 37 |  |  | 20
| 1 | 3 | 39 | 38.91 | 41.12 | 46.26 | 43.3297 | 36.7 | 48 | 48 | 39 | 39 | 38 | 40 | 40 | 40 | 55.466 | 40 | 28 | 42.5272 | 40 | 31.5 | 40 | 35.75 | 38.466 | 31 | 40 | 38 | 44 | 36 | 35 |  |  | 16
| 1 | 4 | 39 | 39.48 | 40.31 | 46.45 | 40.3252 | 37.6 | 45 | 48 | 38 | 40 | 39 | 40 | 40 | 40 | 53.492 | 40 | 27 | 42.1982 | 38 | 31.6 | 40 | 35.14 | 39.096 | 31 | 40 | 39 | 42 | 35 | 35 |  |  | 39.1255
| 1 | 5 | 39 | 39.4 | 40.31 | 45.85 | 39.6344 | 37.3 | 46 | 48 | 38 | 40 | 39 | 40 | 40 | 40 | 51.467 | 40 | 27 | 41.1914 | 41 | 30.7 | 42 | 34.62 | 38.376 | 30 | 40 | 39 | 43 | 36 | 35 |  |  | 79
| 1 | 6 | 39 | 38.96 | 41.4 | 45.72 | 40.9096 | 39.2 | 48 | 47 | 39 | 40 | 39 | 40 | 40 | 40 | 54.027 | 40 | 27 | 41.4477 | 41 | 30.9 | 41 | 35.71 | 38.214 | 29 | 40 | 37 | 41 | 35 | 35
| Mean |  | 38.6667 | 39.01 | 41.0217 | 46.2233 | 40.9827 | 38.5 | 47.3333 | 47 | 38.5 | 39.6667 | 39 | 40 | 40 | 40 | 53.4422 | 40 | 27.1667 | 41.3202 | 39.6667 | 31.5 | 40.3333 | 35.045 | 37.851 | 30.8333 | 40 | 38.3333 | 42.3333 | 35.5 | 35.5
| Median |  | 39 | 39.015 | 40.875 | 46.23 | 40.7385 | 38.4 | 48 | 47.5 | 38.5 | 40 | 39 | 40 | 40 | 40 | 53.5265 | 40 | 27 | 41.3676 | 40 | 31.55 | 40 | 35.205 | 38.295 | 31 | 40 | 38.5 | 42.5 | 35.5 | 35
| Std Dev. |  | 0.5164 | 0.44317 | 0.78871 | 0.41263 | 1.2616 | 1.5375 | 1.5055 | 1.2649 | 0.54772 | 0.5164 | 0.63246 | 0 | 0 | 0 | 1.3416 | 0 | 0.75277 | 1.1373 | 1.3663 | 0.62929 | 1.0328 | 0.74562 | 1.117 | 1.169 | 0 | 0.8165 | 1.2111 | 0.54772 | 0.83666
| Rel.Std.Dev. |  | 1.336% | 1.136% | 1.923% | 0.8927% | 3.078% | 3.994% | 3.181% | 2.691% | 1.423% | 1.302% | 1.622% | 0% | 0% | 0% | 2.51% | 0% | 2.771% | 2.752% | 3.444% | 1.998% | 2.561% | 2.128% | 2.951% | 3.791% | 0% | 2.13% | 2.861% | 1.543% | 2.357%
| PDM3 |  | -1.173% | -0.2952% | 4.846% | 18.14% | 4.747% | -1.599% | 20.98% | 20.13% | -1.599% | 1.383% | -0.3208% | 2.235% | 2.235% | 2.235% | 36.59% | 2.235% | -30.57% | 5.609% | 1.383% | -19.49% | 3.087% | -10.43% | -3.257% | -21.19% | 2.235% | -2.025% | 8.199% | -9.266% | -9.266%
| Z-Score (Absolute) |  | 0.51 | 0.33 | 0.69 | 3.32 | 0.67 | 0.59 | 3.88 | 3.71 | 0.59 | 0 | 0.34 | 0.17 | 0.17 | 0.17 | 6.97 | 0.17 | 6.32 | 0.84 | 0 | 4.13 | 0.34 | 2.34 | 0.92 | 4.47 | 0.17 | 0.67 | 1.35 | 2.11 | 2.11
| Analytical results for Be in OREAS 233b (Certified Value 0.2 ppm) | [32 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.2 | 0.21 | < 1 | 0.16111 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.24 | < 5 | < 1 | 0.2 | 0.2 | 0.2 | < 10 | 0.19 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.18 | 0.09 | 0.3 | 0.24 | 0.16 | [5 empty] | 1
| 1 | 2 | 0.1 | 0.25 | < 1 | 0.15867 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.24 | < 5 | < 1 | 0.21 | 0.21 | 0.2 | < 10 | 0.2 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.18 | < 0.05 | 0.2 | 0.26 | 0.16 | [5 empty] | 21
| 1 | 3 | 0.2 | 0.25 | < 1 | 0.1603 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.22 | < 5 | < 1 | 0.21 | 0.21 | 0.2 | < 10 | 0.2 | < 0.5 | 1.0247 | < 0.5 | < 0.5 | 0.2 | 0.18 | 0.06 | 0.3 | 0.24 | 0.13 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.23 | < 1 | 0.15399 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.23 | < 5 | < 1 | 0.21 | 0.21 | 0.2 | < 10 | 0.2 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.18 | 0.07 | 0.3 | 0.24 | 0.17 | [5 empty] | 0.20267
| 1 | 5 | 0.2 | 0.22 | < 1 | 0.14513 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.2 | < 5 | < 1 | 0.2 | 0.21 | 0.2 | < 10 | 0.2 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.18 | 0.05 | 0.3 | 0.25 | 0.2 | [5 empty] | 80
| 1 | 6 | 0.2 | 0.24 | < 1 | 0.15084 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.22 | < 5 | < 1 | 0.21 | 0.21 | 0.2 | < 10 | 0.2 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.18 | 0.08 | 0.2 | 0.24 | 0.23
| Mean |  | 0.18333 | 0.23333 |  | 0.15501 |  | 0.21667 |  |  | 0.225 |  |  | 0.20667 | 0.20833 | 0.2 |  | 0.19833 |  | 1.0247 |  |  | 0.2 | 0.18 | 0.07 | 0.26667 | 0.245 | 0.175
| Median |  | 0.2 | 0.235 |  | 0.15633 |  | 0.2 |  |  | 0.225 |  |  | 0.21 | 0.21 | 0.2 |  | 0.2 |  | 1.0247 |  |  | 0.2 | 0.18 | 0.07 | 0.3 | 0.24 | 0.165
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.01633; 1.04: 0.0062393; 1.07: 0.040825; 1.11: 0.015166; 1.14: 0.005164; 1.15: 0.0040825; 1.16: 3.0405e-17; 1.18: 0.0040825; 1.27: 3.0405e-17; 1.29: 3.0405e-17; 1.31: 0.015811; 1.32: 0.05164; 1.33: 0.0083666; 1.34: 0.035071
| Rel.Std.Dev. |  | 1.01: 22.27%; 1.02: 6.999%; 1.04: 4.025%; 1.07: 18.84%; 1.11: 6.74%; 1.14: 2.499%; 1.15: 1.96%; 1.16: 1.52e-14%; 1.18: 2.058%; 1.27: 1.52e-14%; 1.29: 1.689e-14%; 1.31: 22.59%; 1.32: 19.36%; 1.33: 3.415%; 1.34: 20.04%
| PDM3 |  | -9.54% | 15.13% |  | -23.52% |  | 6.908% |  |  | 11.02% |  |  | 1.973% | 2.796% | -1.316% |  | -2.138% |  | 405.6% |  |  | -1.316% | -11.18% | -65.46% | 31.58% | 20.89% | -13.65%
| Z-Score (Absolute) |  | NA | 0.22 | 3.86 | 1.41 | 0.04 | NA | 0.04 | 0.04 | 0.34 | 34.39 | 3.86 | 0.62 | 0.6 | 0.73 | 72.56 | 0.75 | 0.04 | 5.19 | 0.04 | 0.04 | NA | 1.03 | 2.82 | NA | 0.04 | 1.11
Results from laboratories 1.01, 1.07, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 233b (Certified Value 0.041 ppm) | [32 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.02 | 0.04 | < 1 | 0.024508 | 0.05 | 0.06 | < 0.05 | 0.05 | < 2 | < 0.1 | 0.04 | 0.04 | 0.04 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.07 | < 0.1 | 0.04378 | 0.04 | 0.087 | 0.06 | 0.04 | 0.05 | 0.03 |  |  |  |  | 1
| 1 | 2 | 0.02 | 0.04 | < 1 | 0.025797 | 0.05 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.03 | 0.04 | 0.04 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.04 | < 0.1 | 0.04451 | 0.04 | 0.052 | 0.07 | 0.05 | 0.05 | 0.03 |  |  |  |  | 22
| 1 | 3 | < 0.02 | 0.04 | < 1 | 0.037534 | 0.04 | 0.06 | 0.05 | 0.05 | < 2 | < 0.1 | 0.04 | 0.04 | 0.04 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.04 | < 0.1 | 0.04802 | 0.04 | 0.043 | 0.06 | 0.05 | 0.04 | 0.03 |  |  |  |  | 16
| 1 | 4 | 0.02 | 0.05 | < 1 | 0.026011 | 0.04 | 0.07 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.04 | 0.04 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.05 | < 0.1 | 0.04411 | 0.03 | 0.032 | 0.06 | 0.05 | 0.05 | 0.03 |  |  |  |  | 0.041038
| 1 | 5 | < 0.02 | 0.04 | < 1 | 0.021315 | 0.04 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.04 | 0.04 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.04 | < 0.1 | 0.04621 | 0.04 | 0.029 | 0.06 | 0.05 | 0.05 | 0.03 |  |  |  |  | 81
| 1 | 6 | < 0.02 | 0.04 | < 1 | 0.029646 | 0.04 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.04 | 0.05 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.04 | < 0.1 | 0.04145 | 0.04 | 0.027 | 0.05 | 0.05 | 0.05 | 0.03
| Mean |  | 0.02 | 0.041667 |  | 0.027468 | 0.043333 | 0.061667 | 0.05 | 0.043333 |  |  | 0.038333 | 0.04 | 0.041667 |  | 0.04 |  |  | 0.03 | 0.046667 |  | 0.04468 | 0.038333 | 0.045 | 0.06 | 0.048333 | 0.048333 | 0.03
| Median |  | 0.02 | 0.04 |  | 0.025904 | 0.04 | 0.06 | 0.05 | 0.04 |  |  | 0.04 | 0.04 | 0.04 |  | 0.04 |  |  | 0.03 | 0.04 |  | 0.04431 | 0.04 | 0.0375 | 0.06 | 0.05 | 0.05 | 0.03
| Std Dev. |  | 0 | 0.0040825 |  | 0.0056142 | 0.005164 | 0.0040825 |  | 0.005164 |  |  | 0.0040825 | 0 | 0.0040825 |  | 0 |  |  | 0 | 0.012111 |  | 0.0022402 | 0.0040825 | 0.022654 | 0.0063246 | 0.0040825 | 0.0040825 | 0
| Rel.Std.Dev. |  | 0% | 9.798% |  | 20.44% | 11.92% | 6.62% |  | 11.92% |  |  | 10.65% | 0% | 9.798% |  | 0% |  |  | 0% | 25.95% |  | 5.014% | 10.65% | 50.34% | 10.54% | 8.447% | 8.447% | 0%
| PDM3 |  | -51.27% | 1.531% |  | -33.07% | 5.592% | 50.27% | 21.84% | 5.592% |  |  | -6.591% | -2.53% | 1.531% |  | -2.53% |  |  | -26.9% | 13.71% |  | 8.874% | -6.591% | 9.654% | 46.2% | 17.78% | 17.78% | -26.9%
| Z-Score (Absolute) |  | 3.03 | 0.17 | 46.19 | 1.6 | 0 | 1.85 | 1.43 | 0 | 96.76 | 0.67 | 0.51 | 0.34 | 0.17 | 501.35 | 0.34 | 248.48 | 1.85 | 1.35 | 0.34 | 0.67 | 0.14 | 0.51 | 0.17 | 1.69 | 0.51 | 0.51 | 1.35
| Analytical results for Ca in OREAS 233b (Certified Value 2.25 wt.%) | [32 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 2.14 | 2.229 | 2.5134 | 2.8484 | 2.4292 | 3.01 | 2.652 | 2.87 | 2.56 | 2.44 | 2.07 | 2.27 | 2.26 | 2.17 | 1.6345 | 2.17 | 2.07 | 2.4769 | 2.32 | 2.06 | 2.43 | 1.905 | 2.3569 | 1.5545 | 2.17 | 1.88 | 2.06 | 2.11 | 1.72 |  |  | 1
| 1 | 2 | 2.23 | 2.225 | 2.4424 | 2.8168 | 2.4045 | 2.97 | 2.666 | 2.9 | 2.55 | 2.4 | 2.05 | 2.32 | 2.33 | 2.2 | 1.6891 | 2.3 | 2.28 | 2.7118 | 2.34 | 2.18 | 2.29 | 2.2156 | 2.3017 | 1.5325 | 2.15 | 2.04 | 2.05 | 2.2 | 1.76 |  |  | 0
| 1 | 3 | 2.3 | 2.25 | 2.3732 | 2.8126 | 2.5903 | 2.81 | 2.694 | 2.95 | 2.53 | 2.38 | 2.03 | 2.3 | 2.31 | 2.2 | 1.7601 | 2.19 | 2.25 | 2.8008 | 2.44 | 2.11 | 2.29 | 1.7251 | 2.4485 | 1.4745 | 2.13 | 1.96 | 2.09 | 2.16 | 1.77 |  |  | 16
| 1 | 4 | 2.31 | 2.282 | 2.3332 | 2.811 | 2.3549 | 2.76 | 2.473 | 2.91 | 2.49 | 2.39 | 2.02 | 2.31 | 2.3 | 2.22 | 1.733 | 2.23 | 2.17 | 2.7529 | 2.28 | 1.99 | 2.22 | 2.1433 | 2.5129 | 1.5295 | 2.14 | 1.9 | 2.06 | 2.14 | 1.7 |  |  | 2.2485
| 1 | 5 | 2.31 | 2.264 | 2.4108 | 2.8274 | 2.3822 | 2.88 | 2.666 | 3 | 2.48 | 2.44 | 1.91 | 2.29 | 2.31 | 2.21 | 1.5815 | 2.27 | 2.09 | 2.6793 | 2.36 | 2.08 | 2.23 | 1.6889 | 2.459 | 1.452 | 2.1 | 1.83 | 2.1 | 2.13 | 1.77 |  |  | 82
| 1 | 6 | 2.24 | 2.216 | 2.5236 | 2.8327 | 2.4245 | 2.99 | 2.694 | 3 | 2.59 | 2.36 | 1.94 | 2.29 | 2.34 | 2.24 | 1.604 | 2.31 | 2.06 | 2.7108 | 2.41 | 2.07 | 2.2 | 2.1132 | 2.4206 | 1.443 | 2.13 | 1.74 | 2.08 | 2.12 | 1.73
| Mean |  | 2.255 | 2.2443 | 2.4328 | 2.8248 | 2.431 | 2.9033 | 2.6408 | 2.9383 | 2.5333 | 2.4017 | 2.0033 | 2.2967 | 2.3083 | 2.2067 | 1.667 | 2.245 | 2.1533 | 2.6888 | 2.3583 | 2.0817 | 2.2767 | 1.9652 | 2.4166 | 1.4977 | 2.1367 | 1.8917 | 2.0733 | 2.1433 | 1.7417
| Median |  | 2.27 | 2.2395 | 2.4266 | 2.8221 | 2.4145 | 2.925 | 2.666 | 2.93 | 2.54 | 2.395 | 2.025 | 2.295 | 2.31 | 2.205 | 1.6618 | 2.25 | 2.13 | 2.7113 | 2.35 | 2.075 | 2.26 | 2.0091 | 2.4346 | 1.502 | 2.135 | 1.89 | 2.07 | 2.135 | 1.745
| Std Dev. |  | 0.066558 | 0.025492 | 0.075891 | 0.014349 | 0.082822 | 0.10309 | 0.08392 | 0.054191 | 0.042269 | 0.032506 | 0.06377 | 0.017512 | 0.027869 | 0.023381 | 0.071906 | 0.057879 | 0.095219 | 0.11194 | 0.058793 | 0.062423 | 0.083825 | 0.22535 | 0.075988 | 0.047047 | 0.023381 | 0.10362 | 0.019664 | 0.03266 | 0.029269
| Rel.Std.Dev. |  | 2.952% | 1.136% | 3.12% | 0.5079% | 3.407% | 3.551% | 3.178% | 1.844% | 1.669% | 1.353% | 3.183% | 0.7625% | 1.207% | 1.06% | 4.313% | 2.578% | 4.422% | 4.163% | 2.493% | 2.999% | 3.682% | 11.47% | 3.144% | 3.141% | 1.094% | 5.478% | 0.9484% | 1.524% | 1.681%
| PDM3 |  | 0.2874% | -0.187% | 8.194% | 25.63% | 8.112% | 29.12% | 17.45% | 30.68% | 12.67% | 6.81% | -10.91% | 2.14% | 2.659% | -1.862% | -25.86% | -0.1574% | -4.234% | 19.58% | 4.883% | -7.421% | 1.251% | -12.6% | 7.474% | -33.39% | -4.975% | -15.87% | -7.792% | -4.679% | -22.54%
| Z-Score (Absolute) |  | 0 | 0.04 | 0.68 | 2.18 | 0.67 | 2.48 | 1.48 | 2.62 | 1.07 | 0.56 | 0.96 | 0.16 | 0.2 | 0.19 | 2.25 | 0.04 | 0.39 | 1.66 | 0.4 | 0.66 | 0.08 | 1.11 | 0.62 | 2.9 | 0.45 | 1.39 | 0.7 | 0.43 | 1.97
| Analytical results for Cd in OREAS 233b (Certified Value 0.54 ppm) | [32 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.58 | 0.532 | < 1 | 0.45555 | 0.60008 | 0.5 | < 1 | 0.51 | 0.59 | 0.7 | 0.5 | 0.53 | 0.53 | 0.49 | < 10 | 0.59 | < 1 | < 1 | 0.49 | 0.48 | 0.56 | 1.57 | 0.5 | 0.51074 | 0.55 | 0.5 | 0.56 | 0.52 | 0.56 |  |  | 1
| 1 | 2 | 0.58 | 0.553 | < 1 | 0.43174 | 0.56974 | 0.59 | < 1 | 0.51 | 0.55 | 0.7 | 0.5 | 0.55 | 0.55 | 0.52 | < 10 | 0.57 | < 1 | < 1 | 0.52 | 0.49 | 0.55 | 1.51 | 0.5 | 0.52813 | 0.56 | 0.5 | 0.6 | 0.53 | 0.58 |  |  | 23
| 1 | 3 | 0.65 | 0.557 | < 1 | 0.4104 | 0.56974 | 0.57 | < 1 | 0.56 | 0.56 | 0.7 | 0.5 | 0.59 | 0.55 | 0.49 | < 10 | 0.58 | < 1 | < 1 | 0.55 | 0.5 | 0.55 | 1.52 | < 0.5 | 0.53115 | 0.53 | 0.51 | 0.61 | 0.54 | 0.59 |  |  | 16
| 1 | 4 | 0.62 | 0.546 | < 1 | 0.42969 | 0.57178 | 0.57 | < 1 | 0.53 | 0.52 | 0.7 | 0.5 | 0.53 | 0.55 | 0.55 | < 10 | 0.62 | < 1 | < 1 | 0.5 | 0.48 | 0.55 | 1.6 | < 0.5 | 0.55489 | 0.53 | 0.51 | 0.58 | 0.55 | 0.59 |  |  | 0.54255
| 1 | 5 | 0.57 | 0.564 | < 1 | 0.44155 | 0.5864 | 0.46 | < 1 | 0.5 | 0.54 | 0.7 | 0.5 | 0.54 | 0.56 | 0.53 | < 10 | 0.64 | < 1 | < 1 | 0.54 | 0.48 | 0.58 | 1.5 | < 0.5 | 0.56772 | 0.54 | 0.51 | 0.59 | 0.53 | 0.57 |  |  | 83
| 1 | 6 | 0.58 | 0.546 | < 1 | 0.43339 | 0.57627 | 0.47 | < 1 | 0.49 | 0.57 | 0.7 | 0.5 | 0.52 | 0.55 | 0.54 | < 10 | 0.62 | < 1 | < 1 | 0.48 | 0.48 | 0.55 | 1.61 | < 0.5 | 0.5345 | 0.56 | 0.52 | 0.57 | 0.55 | 0.55
| Mean |  | 0.59667 | 0.54967 |  | 0.43372 | 0.579 | 0.52667 |  | 0.51667 | 0.555 | 0.7 | 0.5 | 0.54333 | 0.54833 | 0.52 |  | 0.60333 |  |  | 0.51333 | 0.485 | 0.55667 | 1.5517 | 0.5 | 0.53785 | 0.545 | 0.50833 | 0.585 | 0.53667 | 0.57333
| Median |  | 0.58 | 0.5495 |  | 0.43256 | 0.57403 | 0.535 |  | 0.51 | 0.555 | 0.7 | 0.5 | 0.535 | 0.55 | 0.525 |  | 0.605 |  |  | 0.51 | 0.48 | 0.55 | 1.545 | 0.5 | 0.53282 | 0.545 | 0.51 | 0.585 | 0.535 | 0.575
| Std Dev. |  | 0.031411 | 0.011039 |  | 0.014848 | 0.012084 | 0.056804 |  | 0.025033 | 0.02429 | 1.2162e-16 | 0 | 0.025033 | 0.0098319 | 0.025298 |  | 0.027325 |  |  | 0.028048 | 0.0083666 | 0.012111 | 0.047924 | 0 | 0.020338 | 0.013784 | 0.0075277 | 0.018708 | 0.012111 | 0.01633
| Rel.Std.Dev. |  | 5.264% | 2.008% |  | 3.423% | 2.087% | 10.79% |  | 4.845% | 4.377% | 1.737e-14% | 0% | 4.607% | 1.793% | 4.865% |  | 4.529% |  |  | 5.464% | 1.725% | 2.176% | 3.089% | 0% | 3.781% | 2.529% | 1.481% | 3.198% | 2.257% | 2.848%
| PDM3 |  | 9.975% | 1.312% |  | -20.06% | 6.719% | -2.927% |  | -4.77% | 2.295% | 29.02% | -7.842% | 0.1446% | 1.066% | -4.156% |  | 11.2% |  |  | -5.385% | -10.61% | 2.602% | 186% | -7.842% | -0.8651% | 0.4518% | -6.306% | 7.824% | -1.084% | 5.674%
| Z-Score (Absolute) |  | 1.12 | 0.22 | 0.72 | 1.98 | 0.78 | 0.21 | 0.72 | 0.4 | 0.33 | NA | NA | 0.1 | 0.2 | 0.34 | 84.81 | 1.24 | 0.72 | 0.72 | 0.47 | 1 | 0.36 | 19.27 | 3.89 | 0 | 0.14 | 0.56 | 0.9 | 0.02 | 0.67
Results from laboratories 1.12 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 233b (Certified Value 9.28 ppm) | [32 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | 8.58 | 9.326 | 10.045 | 8.2697 | 9.6668 | 10.2 | < 1 | 10.8 | 9.388 | 10 | 8.55 | 9.31 | 8.85 | 9.25 | < 10 | 8.95 | 9.81 | 9 | 9.6 | 9.92 | 9.34 | 9.63 | 9.62 | [8 empty] | 1
| 1 | 2 | 8.76 | 9.289 | 8.987 | 8.5341 | 9.6547 | 9.91 | < 1 | 10.9 | 9.154 | 11 | 8.6 | 9.34 | 9.09 | 9.29 | < 10 | 9.12 | 9.69 | 9 | 9.68 | 10.1 | 9.78 | 9.87 | 9.72 | [8 empty] | 24
| 1 | 3 | 8.79 | 9.501 | 8.392 | 8.4285 | 11.2822 | 9.82 | < 1 | 11.3 | 9.187 | 10 | 8.75 | 9.23 | 9.21 | 9.16 | < 10 | 9.14 | 9.76 | 8.5 | 9.43 | 10.1 | 9.42 | 9.85 | 9.55 | [8 empty] | 16
| 1 | 4 | 8.76 | 9.482 | 8.373 | 8.3664 | 9.4894 | 9.61 | < 1 | 11.5 | 8.857 | 10 | 8.39 | 9.47 | 9.33 | 9.17 | < 10 | 8.58 | 9.98 | 8.5 | 9.6 | 10.1 | 9.33 | 9.67 | 9.47 | [8 empty] | 9.2757
| 1 | 5 | 8.75 | 9.333 | 8.62 | 8.1092 | 9.1643 | 10 | < 1 | 11.4 | 9.202 | 10 | 8.66 | 9.46 | 9.18 | 9.7 | < 10 | 8.47 | 9.96 | 8.5 | 9.65 | 10.3 | 9.51 | 9.53 | 9.47 | [8 empty] | 84
| 1 | 6 | 8.78 | 9.303 | 9.044 | 8.4466 | 9.4688 | 10.1 | < 1 | 10.9 | 9.425 | 10 | 8.43 | 9.52 | 9.58 | 9.63 | < 10 | 9.26 | 9.76 | 8.5 | 9.64 | 9.97 | 9.06 | 9.61 | 9.23
| Mean |  | 8.7367 | 9.3723 | 8.9102 | 8.3591 | 9.7877 | 9.94 |  | 11.1333 | 9.2022 | 10.1667 | 8.5633 | 9.3883 | 9.2067 | 9.3667 |  | 8.92 | 9.8267 | 8.6667 | 9.6 | 10.0817 | 9.4067 | 9.6933 | 9.51
| Median |  | 8.76 | 9.3295 | 8.8035 | 8.3975 | 9.572 | 9.955 |  | 11.1 | 9.1945 | 10 | 8.575 | 9.4 | 9.195 | 9.27 |  | 9.035 | 9.785 | 8.5 | 9.62 | 10.1 | 9.38 | 9.65 | 9.51
| Std Dev. |  | 0.078145 | 0.093837 | 0.62496 | 0.15074 | 0.75435 | 0.21024 |  | 0.30111 | 0.20292 | 0.40825 | 0.13677 | 0.11197 | 0.24337 | 0.23721 |  | 0.32342 | 0.11759 | 0.2582 | 0.088769 | 0.13212 | 0.23695 | 0.13706 | 0.16697
| Rel.Std.Dev. |  | 0.8945% | 1.001% | 7.014% | 1.803% | 7.707% | 2.115% |  | 2.705% | 2.205% | 4.016% | 1.597% | 1.193% | 2.643% | 2.532% |  | 3.626% | 1.197% | 2.979% | 0.9247% | 1.311% | 2.519% | 1.414% | 1.756%
| PDM3 |  | -5.812% | 1.041% | -3.941% | -9.882% | 5.519% | 7.161% |  | 20.03% | -0.7933% | 9.605% | -7.68% | 1.214% | -0.7447% | 0.9802% |  | -3.835% | 5.939% | -6.566% | 3.496% | 8.688% | 1.411% | 4.502% | 2.525%
| Z-Score (Absolute) |  | 0.94 | 0 | 0.68 | 1.5 | 0.62 | 0.85 | 13.19 | 2.62 | 0.25 | NA | 1.2 | 0.03 | 0.24 | 0 | 6.5 | 0.67 | 0.68 | 1.05 | 0.34 | 1.06 | 0.06 | 0.48 | 0.21
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 233b (Certified Value 31.6 ppm) | [32 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 29.7 | 31.22 | 29.6 | 25.692 | 32.3032 | 32.5 | 30.8 | 35.5 | 32.4 | 29 | 32.6 | 34.3 | 31.7 | 30.4 | 23.801 | 30.8 | 29 | 34.0226 | 33.8 | 30 | 33 | 28.29 | 28 | 33.6701 | 30.2 | 34.4 | 32 | 31.4 | 31.5 |  |  | 1
| 1 | 2 | 30.6 | 31.02 | 28.49 | 25.4809 | 32.1811 | 32.6 | 31.8 | 34.1 | 32.7 | 28 | 35.7 | 34 | 32 | 31.6 | 25.402 | 30.9 | 30 | 34.7827 | 32.1 | 31.9 | 32 | 27.49 | 28 | 34.2657 | 31 | 35.8 | 33.8 | 31.5 | 31.6 |  |  | 25
| 1 | 3 | 30.4 | 31.69 | 27.84 | 25.526 | 34.8041 | 32.2 | 30.8 | 34.3 | 32.5 | 28 | 31.8 | 34.7 | 31.8 | 31.5 | 26.105 | 30.5 | 30 | 34.889 | 33.6 | 31.1 | 32 | 29.63 | 27 | 32.4175 | 30.8 | 35.5 | 33.6 | 31.5 | 31.1 |  |  | 16
| 1 | 4 | 30.1 | 31.32 | 27.55 | 25.5795 | 32.2731 | 31.7 | 32.2 | 34.7 | 32.1 | 28 | 32.2 | 32.8 | 32.2 | 31.8 | 24.381 | 30.6 | 29 | 34.78 | 31.4 | 30.3 | 32 | 28.81 | 28 | 33.609 | 30.8 | 36.3 | 33.5 | 31.1 | 30.9 |  |  | 31.5941
| 1 | 5 | 30.1 | 31.09 | 28.29 | 25.3146 | 32.3811 | 32.4 | 31.5 | 34.8 | 32.1 | 29 | 32.2 | 34 | 31.9 | 31.2 | 23.817 | 32.5 | 29 | 34.3782 | 32.7 | 30.4 | 32 | 27.38 | 27 | 33.8256 | 30.7 | 35.7 | 32.8 | 31.1 | 31 |  |  | 85
| 1 | 6 | 30.2 | 30.81 | 28.52 | 25.0565 | 32.4883 | 33.1 | 32 | 34 | 32.8 | 29 | 32.2 | 33.7 | 31.9 | 32.8 | 25.608 | 31.9 | 29 | 35.6392 | 31.5 | 29.9 | 32 | 28.72 | 26 | 33.5553 | 30.5 | 35.3 | 32 | 31.2 | 30.5
| Mean |  | 30.1833 | 31.1917 | 28.3817 | 25.4416 | 32.7385 | 32.4167 | 31.5167 | 34.5667 | 32.4333 | 28.5 | 32.7833 | 33.9167 | 31.9167 | 31.55 | 24.8523 | 31.2 | 29.3333 | 34.7486 | 32.5167 | 30.6 | 32.1667 | 28.3867 | 27.3333 | 33.5572 | 30.6667 | 35.5 | 32.95 | 31.3 | 31.1
| Median |  | 30.15 | 31.155 | 28.39 | 25.5035 | 32.3422 | 32.45 | 31.65 | 34.5 | 32.45 | 28.5 | 32.2 | 34 | 31.9 | 31.55 | 24.8915 | 30.85 | 29 | 34.7813 | 32.4 | 30.35 | 32 | 28.505 | 27.5 | 33.6395 | 30.75 | 35.6 | 33.15 | 31.3 | 31.05
| Std Dev. |  | 0.30605 | 0.30036 | 0.70833 | 0.22577 | 1.0172 | 0.46224 | 0.60139 | 0.55737 | 0.29439 | 0.54772 | 1.4511 | 0.64317 | 0.17224 | 0.78422 | 0.98404 | 0.80994 | 0.5164 | 0.54394 | 1.0304 | 0.7642 | 0.40825 | 0.85591 | 0.8165 | 0.61454 | 0.28048 | 0.63561 | 0.80932 | 0.18974 | 0.40497
| Rel.Std.Dev. |  | 1.014% | 0.963% | 2.496% | 0.8874% | 3.107% | 1.426% | 1.908% | 1.612% | 0.9077% | 1.922% | 4.426% | 1.896% | 0.5397% | 2.486% | 3.96% | 2.596% | 1.76% | 1.565% | 3.169% | 2.497% | 1.269% | 3.015% | 2.987% | 1.831% | 0.9146% | 1.79% | 2.456% | 0.6062% | 1.302%
| PDM3 |  | -4.465% | -1.274% | -10.17% | -19.47% | 3.622% | 2.604% | -0.2451% | 9.409% | 2.656% | -9.793% | 3.764% | 7.351% | 1.021% | -0.1396% | -21.34% | -1.247% | -7.156% | 9.984% | 2.92% | -3.146% | 1.812% | -10.15% | -13.49% | 6.214% | -2.935% | 12.36% | 4.292% | -0.9309% | -1.564%
| Z-Score (Absolute) |  | 0.71 | 0.17 | 1.67 | 3.23 | 0.65 | 0.48 | 0 | 1.62 | 0.49 | 1.61 | 0.67 | 1.28 | 0.21 | 0.02 | 3.55 | 0.17 | 1.16 | 1.72 | 0.53 | 0.49 | 0.35 | 1.67 | 2.23 | 1.09 | 0.45 | 2.12 | 0.76 | 0.12 | 0.22
| Analytical results for Cr in OREAS 233b (Certified Value 28.4 ppm) | [32 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 30 | 28.7 | 30.44 | 29.1177 | 35 | 32 | 35 | 33 | 28 | 32 | 28 | 27 | 26 | 22.916 | 27 | 24 | 28.3242 | 31 | 27.8 | 29 | 31.07 | 28.2255 | 25 | 27 | 26.3 | 29 | 25 | 26.7 |  |  |  | 1
| 1 | 2 | 30 | 28.4 | 29.99 | 29.0128 | 34 | 33 | 34 | 31 | 27 | 31 | 28 | 28 | 27 | 23.936 | 27 | 26 | 28.824 | 32 | 29 | 29 | 26.7 | 28.1604 | 25 | 27 | 28.1 | 29 | 26 | 26.7 |  |  |  | 26
| 1 | 3 | 30 | 29.2 | 29.87 | 31.5512 | 33 | 34 | 35 | 32 | 26 | 31 | 27 | 28 | 27 | 24.763 | 27 | 25 | 30.162 | 32 | 28.1 | 28 | 28.09 | 29.5275 | 25 | 26 | 26.3 | 28 | 26 | 27.1 |  |  |  | 16
| 1 | 4 | 31 | 28.7 | 29.98 | 28.712 | 33 | 32 | 37 | 31 | 28 | 31 | 27 | 28 | 27 | 23.995 | 27 | 25 | 29.7064 | 30 | 27.7 | 28 | 30.06 | 29.6484 | 25 | 26 | 27.4 | 29 | 25 | 26.4 |  |  |  | 28.3667
| 1 | 5 | 31 | 28.8 | 30.13 | 28.8587 | 34 | 33 | 35 | 30 | 28 | 31 | 27 | 28 | 27 | 22.498 | 28 | 24 | 28.9009 | 30 | 27.7 | 28 | 27.89 | 29.2392 | 25 | 26 | 26.8 | 28 | 26 | 25.8 |  |  |  | 86
| 1 | 6 | 31 | 28.5 | 30.47 | 29.0451 | 35 | 32 | 35 | 31 | 28 | 31 | 27 | 28 | 27 | 23.954 | 28 | 24 | 30.047 | 31 | 28.1 | 28 | 29.63 | 28.3836 | 20 | 27 | 26.7 | 29 | 25 | 25.3
| Mean |  | 30.5 | 28.7167 | 30.1467 | 29.3829 | 34 | 32.6667 | 35.1667 | 31.3333 | 27.5 | 31.1667 | 27.3333 | 27.8333 | 26.8333 | 23.677 | 27.3333 | 24.6667 | 29.3274 | 31 | 28.0667 | 28.3333 | 28.9067 | 28.8641 | 24.1667 | 26.5 | 26.9333 | 28.6667 | 25.5 | 26.3333
| Median |  | 30.5 | 28.7 | 30.06 | 29.0289 | 34 | 32.5 | 35 | 31 | 28 | 31 | 27 | 28 | 27 | 23.945 | 27 | 24.5 | 29.3037 | 31 | 27.95 | 28 | 28.86 | 28.8114 | 25 | 26.5 | 26.75 | 29 | 25.5 | 26.55
| Std Dev. |  | 0.54772 | 0.27869 | 0.25288 | 1.0722 | 0.89443 | 0.8165 | 0.98319 | 1.0328 | 0.83666 | 0.40825 | 0.5164 | 0.40825 | 0.40825 | 0.82383 | 0.5164 | 0.8165 | 0.74829 | 0.89443 | 0.49261 | 0.5164 | 1.6188 | 0.68261 | 2.0412 | 0.54772 | 0.70048 | 0.5164 | 0.54772 | 0.66533
| Rel.Std.Dev. |  | 1.796% | 0.9705% | 0.8388% | 3.649% | 2.631% | 2.499% | 2.796% | 3.296% | 3.042% | 1.31% | 1.889% | 1.467% | 1.521% | 3.479% | 1.889% | 3.31% | 2.552% | 2.885% | 1.755% | 1.823% | 5.6% | 2.365% | 8.447% | 2.067% | 2.601% | 1.801% | 2.148% | 2.527%
| PDM3 |  | 7.52% | 1.234% | 6.275% | 3.582% | 19.86% | 15.16% | 23.97% | 10.46% | -3.055% | 9.871% | -3.643% | -1.88% | -5.406% | -16.53% | -3.643% | -13.04% | 3.387% | 9.283% | -1.058% | -0.1177% | 1.903% | 1.753% | -14.81% | -6.581% | -5.053% | 1.057% | -10.11% | -7.168%
| Z-Score (Absolute) |  | 0.81 | 0.09 | 0.67 | 0.36 | 2.24 | 1.7 | 2.71 | 1.15 | 0.41 | 1.09 | 0.47 | 0.27 | 0.68 | 1.96 | 0.47 | 1.56 | 0.34 | 1.02 | 0.18 | 0.07 | 0.17 | 0.15 | 1.76 | 0.81 | 0.64 | 0.07 | 1.22 | 0.88
| Analytical results for Cs in OREAS 233b (Certified Value 0.65 ppm) | [32 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.7 | 0.65 | 0.54573 | 0.58 | 0.65 | 0.68 | 0.65 | 0.8 | 0.63 | 0.64 | 0.64 | 0.66 | 0.63 | 0.5 | 0.65 | 0.59 | 0.69 | 0.69 | 0.68 | [12 empty] | 1
| 1 | 2 | 0.7 | 0.64 | 0.55613 | 0.58 | 0.63 | 0.61 | 0.66 | 0.8 | 0.63 | 0.65 | 0.67 | 0.68 | 0.61 | 0.5 | 0.64 | 0.6 | 0.73 | 0.69 | 0.73 | [12 empty] | 27
| 1 | 3 | 0.73 | 0.67 | 0.57022 | 0.56 | 0.61 | 0.7 | 0.66 | 0.7 | 0.64 | 0.65 | 0.67 | 0.65 | 0.61 | 0.5 | 0.65 | 0.58 | 0.72 | 0.69 | 0.68 | [12 empty] | 16
| 1 | 4 | 0.72 | 0.66 | 0.52746 | 0.57 | 0.63 | 0.78 | 0.63 | 0.8 | 0.61 | 0.66 | 0.68 | 0.65 | 0.59 | 0.5 | 0.66 | 0.58 | 0.73 | 0.68 | 0.67 | [12 empty] | 0.65227
| 1 | 5 | 0.7 | 0.64 | 0.5447 | 0.57 | 0.62 | 0.68 | 0.65 | 0.7 | 0.65 | 0.65 | 0.66 | 0.69 | 0.64 | 0.5 | 0.64 | 0.56 | 0.71 | 0.68 | 0.68 | [12 empty] | 87
| 1 | 6 | 0.69 | 0.66 | 0.5546 | 0.55 | 0.63 | 0.69 | 0.67 | 0.8 | 0.64 | 0.66 | 0.69 | 0.68 | 0.6 | 0.5 | 0.65 | 0.56 | 0.69 | 0.68 | 0.68
| Mean |  | 0.70667 | 0.65333 | 0.54981 | 0.56833 | 0.62833 | 0.69 | 0.65333 | 0.76667 | 0.63333 | 0.65167 | 0.66833 | 0.66833 | 0.61333 | 0.5 | 0.64833 | 0.57833 | 0.71167 | 0.685 | 0.68667
| Median |  | 0.7 | 0.655 | 0.55016 | 0.57 | 0.63 | 0.685 | 0.655 | 0.8 | 0.635 | 0.65 | 0.67 | 0.67 | 0.61 | 0.5 | 0.65 | 0.58 | 0.715 | 0.685 | 0.68
| Std Dev. |  | 0.015055 | 0.012111 | 0.014295 | 0.01169 | 0.013292 | 0.054406 | 0.013663 | 0.05164 | 0.013663 | 0.0075277 | 0.017224 | 0.017224 | 0.018619 | 0 | 0.0075277 | 0.016021 | 0.018348 | 0.0054772 | 0.021602
| Rel.Std.Dev. |  | 2.13% | 1.854% | 2.6% | 2.057% | 2.115% | 7.885% | 2.091% | 6.736% | 2.157% | 1.155% | 2.577% | 2.577% | 3.036% | 0% | 1.161% | 2.77% | 2.578% | 0.7996% | 3.146%
| PDM3 |  | 8.339% | 0.1629% | -15.71% | -12.87% | -3.67% | 5.784% | 0.1629% | 17.54% | -2.903% | -0.09263% | 2.463% | 2.463% | -5.97% | -23.34% | -0.6037% | -11.34% | 9.106% | 5.018% | 5.273%
| Z-Score (Absolute) |  | 1.14 | 0 | 2.2 | 1.81 | 0.53 | 0.78 | 0 | NA | 0.43 | 0.04 | 0.32 | 0.32 | 0.85 | NA | 0.11 | 1.6 | 1.24 | 0.67 | 0.71
Results from laboratories 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 233b (Certified Value 160 ppm) | [32 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 176.5 | 164.86 | 150.82 | 173.47 | 165.6054 | 160 | 169 | 152 | 163.8 | 158 | 155.3 | 164.5 | 161 | 158 | 155.187 | 160.5 | 167 | 162.5689 | 157 | 148.89 | 166 | 161.93 | 154.4588 | 155 | 161.8667 | 162.5 | 146.558 | 163 | 159.1 | 151.7 | 160 | 1
| 1 | 2 | 177.1 | 163.07 | 151.23 | 176.48 | 165.1719 | 159 | 164 | 153 | 164.1 | 159 | 164.9 | 163 | 164.5 | 164 | 158.105 | 159.5 | 176 | 163.56 | 161 | 156.08 | 164 | 160.64 | 153.6676 | 156 | 160.2563 | 160.5 | 140.789 | 169 | 164.1 | 154.9 | 161 | 28
| 1 | 3 | 178.6 | 165.88 | 150.13 | 171.44 | 178.9308 | 153 | 168 | 153 | 161.5 | 157 | 152.8 | 163.5 | 162 | 164.5 | 161.812 | 155.5 | 173 | 168.9072 | 162 | 153.13 | 154 | 160.82 | 153.5392 | 152 | 162.7199 | 162 | 140.946 | 161 | 165 | 153 | 162 | 16
| 1 | 4 | 175 | 165.87 | 152.03 | 170.94 | 165.553 | 147 | 167 | 153 | 158.9 | 157 | 153.5 | 161.5 | 161.5 | 162.5 | 157.925 | 160 | 168 | 163.6704 | 154 | 149.7 | 159 | 163.84 | 156.5576 | 155 | 165.5799 | 159.5 | 131.04 | 169 | 165.1 | 152 | 159 | 160.2923
| 1 | 5 | 176.9 | 164.07 | 146.75 | 169.11 | 163.8363 | 152 | 170 | 152 | 160.7 | 160 | 155.2 | 165 | 158.5 | 163 | 152.313 | 167 | 168 | 162.2667 | 155 | 149.34 | 171 | 162.8 | 156.7848 | 155 | 164.1044 | 159.5 | 135.167 | 165 | 161 | 153.4 | 161 | 88
| 1 | 6 | 180.9 | 161.66 | 151.09 | 172.06 | 166.3182 | 156 | 166 | 151 | 160.3 | 161 | 155 | 161.5 | 164 | 165.5 | 160.719 | 165 | 168 | 166.8961 | 154 | 153.81 | 160 | 162.04 | 152.4904 | 153 | 159.776 | 159 | 133.644 | 164 | 161.9 | 153.5 | 158
| Mean |  | 177.5 | 164.235 | 150.3417 | 172.25 | 167.5693 | 154.5 | 167.3333 | 152.3333 | 161.55 | 158.6667 | 156.1167 | 163.1667 | 161.9167 | 162.9167 | 157.6768 | 161.25 | 170 | 164.6449 | 157.1667 | 151.825 | 162.3333 | 162.0117 | 154.5831 | 154.3333 | 162.3839 | 160.5 | 138.024 | 165.1667 | 162.7 | 153.0833 | 160.1667
| Median |  | 177 | 164.465 | 150.955 | 171.75 | 165.5792 | 154.5 | 167.5 | 152.5 | 161.1 | 158.5 | 155.1 | 163.25 | 161.75 | 163.5 | 158.015 | 160.25 | 168 | 163.6152 | 156 | 151.415 | 162 | 161.985 | 154.0632 | 155 | 162.2933 | 160 | 137.978 | 164.5 | 163 | 153.2 | 160.5
| Std Dev. |  | 2.0268 | 1.6604 | 1.864 | 2.5159 | 5.6259 | 4.8477 | 2.1602 | 0.8165 | 2.0433 | 1.633 | 4.4215 | 1.472 | 2.1775 | 2.6347 | 3.5091 | 4.1322 | 3.6332 | 2.6581 | 3.5449 | 2.9344 | 5.9554 | 1.2061 | 1.736 | 1.5055 | 2.2308 | 1.4491 | 5.7469 | 3.2506 | 2.429 | 1.1548 | 1.472
| Rel.Std.Dev. |  | 1.142% | 1.011% | 1.24% | 1.461% | 3.357% | 3.138% | 1.291% | 0.536% | 1.265% | 1.029% | 2.832% | 0.9021% | 1.345% | 1.617% | 2.226% | 2.563% | 2.137% | 1.614% | 2.256% | 1.933% | 3.669% | 0.7444% | 1.123% | 0.9755% | 1.374% | 0.9029% | 4.164% | 1.968% | 1.493% | 0.7544% | 0.919%
| PDM3 |  | 10.74% | 2.46% | -6.208% | 7.46% | 4.54% | -3.614% | 4.393% | -4.965% | 0.7847% | -1.014% | -2.605% | 1.793% | 1.013% | 1.637% | -1.632% | 0.5975% | 6.056% | 2.715% | -1.95% | -5.282% | 1.273% | 1.073% | -3.562% | -3.718% | 1.305% | 0.1296% | -13.89% | 3.041% | 1.502% | -4.497% | -0.07835%
| Z-Score (Absolute) |  | 2.78 | 0.47 | 1.95 | 1.86 | 1.05 | 1.23 | 1.01 | 1.6 | 0 | 0.5 | 0.95 | 0.28 | 0.06 | 0.24 | 0.67 | 0.05 | 1.47 | 0.54 | 0.76 | 1.69 | 0.14 | 0.08 | 1.21 | 1.26 | 0.15 | 0.18 | 4.1 | 0.63 | 0.2 | 1.47 | 0.24
| Analytical results for Dy in OREAS 233b (Certified Value 2.06 ppm) | [32 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.27 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.07: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g | 2
| 1 | 1 | 1.01: 2.073; 1.02: 1.997; 1.03: 1.8586; 1.04: 2.3; 1.05: 2.46; 1.07: 1.8 | 1
| 1 | 2 | 1.01: 2.101; 1.02: 1.987; 1.03: 1.9017; 1.04: 2.4; 1.05: 2.52; 1.07: 1.7 | 29
| 1 | 3 | 1.01: 2.117; 1.02: 1.823; 1.03: 1.9154; 1.04: 2.3; 1.05: 2.43; 1.07: 1.6 | 16
| 1 | 4 | 1.01: 2.122; 1.02: 1.825; 1.03: 1.8605; 1.04: 2.3; 1.05: 2.56; 1.07: 1.6 | 2.0555
| 1 | 5 | 1.01: 2.132; 1.02: 1.937; 1.03: 1.8151; 1.04: 2.3; 1.05: 2.36; 1.07: 1.5 | 89
| 1 | 6 | 1.01: 2.086; 1.02: 2; 1.03: 1.8657; 1.04: 2.3; 1.05: 2.55; 1.07: 1.6
| Mean |  | 1.01: 2.1052; 1.02: 1.9282; 1.03: 1.8695; 1.04: 2.3167; 1.05: 2.48; 1.07: 1.6333
| Median |  | 1.01: 2.109; 1.02: 1.962; 1.03: 1.8631; 1.04: 2.3; 1.05: 2.49; 1.07: 1.6
| Std Dev. |  | 1.01: 0.02264; 1.02: 0.083834; 1.03: 0.035548; 1.04: 0.040825; 1.05: 0.077717; 1.07: 0.10328
| Rel.Std.Dev. |  | 1.01: 1.075%; 1.02: 4.348%; 1.03: 1.901%; 1.04: 1.762%; 1.05: 3.134%; 1.07: 6.323%
| PDM3 |  | 1.01: 2.418%; 1.02: -6.193%; 1.03: -9.048%; 1.04: 12.71%; 1.05: 20.65%; 1.07: -20.54%
| Z-Score (Absolute) |  | 1.01: 0.27; 1.02: 0.27; 1.03: 0.44; 1.04: 0.9; 1.05: 1.4; 1.07: 1.16
| Analytical results for Er in OREAS 233b (Certified Value 1.19 ppm) | [32 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.27 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.07: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g | 2
| 1 | 1 | 1.01: 1.215; 1.02: 1.145; 1.03: 1.117; 1.04: 1.4; 1.05: 1.43; 1.07: 0.95 | 1
| 1 | 2 | 1.01: 1.204; 1.02: 1.144; 1.03: 1.1505; 1.04: 1.4; 1.05: 1.41; 1.07: 0.95 | 30
| 1 | 3 | 1.01: 1.207; 1.02: 1.033; 1.03: 1.158; 1.04: 1.3; 1.05: 1.4; 1.07: 0.9 | 16
| 1 | 4 | 1.01: 1.226; 1.02: 1.048; 1.03: 1.1393; 1.04: 1.3; 1.05: 1.45; 1.07: 0.85 | 1.19
| 1 | 5 | 1.01: 1.231; 1.02: 1.095; 1.03: 1.1446; 1.04: 1.4; 1.05: 1.42; 1.07: 0.85 | 90
| 1 | 6 | 1.01: 1.196; 1.02: 1.158; 1.03: 1.1104; 1.04: 1.4; 1.05: 1.46; 1.07: 0.85
| Mean |  | 1.01: 1.2132; 1.02: 1.1038; 1.03: 1.1366; 1.04: 1.3667; 1.05: 1.4283; 1.07: 0.89167
| Median |  | 1.01: 1.211; 1.02: 1.1195; 1.03: 1.1419; 1.04: 1.4; 1.05: 1.425; 1.07: 0.875
| Std Dev. |  | 1.01: 0.013438; 1.02: 0.053768; 1.03: 0.018928; 1.04: 0.05164; 1.05: 0.023166; 1.07: 0.04916
| Rel.Std.Dev. |  | 1.01: 1.108%; 1.02: 4.871%; 1.03: 1.665%; 1.04: 3.779%; 1.05: 1.622%; 1.07: 5.513%
| PDM3 |  | 1.01: 1.943%; 1.02: -7.245%; 1.03: -4.489%; 1.04: 14.84%; 1.05: 20.02%; 1.07: -25.07%
| Z-Score (Absolute) |  | 1.01: 0.2; 1.02: 0.36; 1.03: 0.2; 1.04: 0.98; 1.05: 1.3; 1.07: 1.45
| Analytical results for Eu in OREAS 233b (Certified Value 0.46 ppm) | [32 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.27 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.07: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g | 2
| 1 | 1 | 1.01: 0.444; 1.02: 0.38; 1.03: 0.49199; 1.04: 0.5; 1.05: 0.51; 1.07: 0.3 | 1
| 1 | 2 | 1.01: 0.432; 1.02: 0.375; 1.03: 0.51333; 1.04: 0.5; 1.05: 0.53; 1.07: 0.3 | 31
| 1 | 3 | 1.01: 0.446; 1.02: 0.355; 1.03: 0.49154; 1.04: 0.5; 1.05: 0.5; 1.07: 0.3 | 16
| 1 | 4 | 1.01: 0.451; 1.02: 0.343; 1.03: 0.48976; 1.04: 0.5; 1.05: 0.53; 1.07: 0.3 | 0.46473
| 1 | 5 | 1.01: 0.453; 1.02: 0.363; 1.03: 0.48383; 1.04: 0.5; 1.05: 0.5; 1.07: 0.3 | 91
| 1 | 6 | 1.01: 0.437; 1.02: 0.391; 1.03: 0.50274; 1.04: 0.5; 1.05: 0.55; 1.07: 0.3
| Mean |  | 1.01: 0.44383; 1.02: 0.36783; 1.03: 0.49553; 1.04: 0.5; 1.05: 0.52; 1.07: 0.3
| Median |  | 1.01: 0.445; 1.02: 0.369; 1.03: 0.49177; 1.04: 0.5; 1.05: 0.52; 1.07: 0.3
| Std Dev. |  | 1.01: 0.008085; 1.02: 0.017555; 1.03: 0.010654; 1.04: 0; 1.05: 0.02; 1.07: 0
| Rel.Std.Dev. |  | 1.01: 1.822%; 1.02: 4.772%; 1.03: 2.15%; 1.04: 0%; 1.05: 3.846%; 1.07: 0%
| PDM3 |  | 1.01: -4.496%; 1.02: -20.85%; 1.03: 6.629%; 1.04: 7.59%; 1.05: 11.89%; 1.07: -35.45%
| Z-Score (Absolute) |  | 1.01: 0.43; 1.02: 1.7; 1.03: 0.43; 1.04: 0.51; 1.05: 0.84; 1.07: 2.84
| Analytical results for Fe in OREAS 233b (Certified Value 5.74 wt.%) | [32 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 5.79 | 5.643 | 3.9965 | 5.5004 | 5.8793 | 5.91 | 6.022 | 6.1 | 5.77 | 5.82 | 5.55 | 5.62 | 5.62 | 5.67 | 5.4822 | 5.78 | 5.68 | > 5 | 6.16 | 5.08 | 5.67 | 6.77 | 5.6262 | 5.048 | 5.49 | 5.5 | 5.77 | 5.57 | 5.67 |  |  | 1
| 1 | 2 | 5.9 | 5.638 | 3.9225 | 5.5091 | 5.8373 | 5.94 | 5.931 | 6.05 | 5.7 | 5.89 | 5.5 | 5.63 | 5.74 | 5.67 | 5.7308 | 5.71 | 5.97 | > 5 | 6.05 | 5.41 | 5.71 | 7.59 | 5.6977 | 5.025 | 5.45 | 5.66 | 5.69 | 5.67 | 5.62 |  |  | 0
| 1 | 3 | 5.9 | 5.776 | 3.9052 | 5.589 | 6.2529 | 5.88 | 6.022 | 6.15 | 5.78 | 5.82 | 5.5 | 5.57 | 5.69 | 5.7 | 5.9168 | 5.64 | 5.87 | > 5 | 5.78 | 5.23 | 5.54 | 5.82 | 5.9873 | 4.942 | 5.49 | 5.5 | 5.66 | 5.6 | 5.5 |  |  | 16
| 1 | 4 | 5.86 | 5.673 | 3.8983 | 5.5054 | 5.8561 | 5.87 | 5.875 | 6.21 | 5.55 | 5.83 | 5.47 | 5.6 | 5.72 | 5.74 | 5.7227 | 5.76 | 5.73 | > 5 | 5.88 | 5.09 | 5.62 | 7.54 | 5.9958 | 5.0535 | 5.44 | 5.6 | 5.68 | 5.55 | 5.59 |  |  | 5.7387
| 1 | 5 | 5.9 | 5.641 | 3.9492 | 5.5387 | 5.8049 | 5.95 | 5.959 | 6.21 | 5.68 | 5.85 | 5.45 | 5.66 | 5.69 | 5.77 | 5.4937 | 5.98 | 5.7 | > 5 | 6.2 | 5.07 | 5.89 | 6.02 | 5.94 | 4.8675 | 5.45 | 5.6 | 5.79 | 5.64 | 5.63 |  |  | 92
| 1 | 6 | 5.91 | 5.642 | 3.9498 | 5.5491 | 5.916 | 5.99 | 5.98 | 6.24 | 5.73 | 5.82 | 5.44 | 5.56 | 5.71 | 5.79 | 5.5098 | 5.93 | 5.66 | > 5 | 5.89 | 5.16 | 5.64 | 7.43 | 5.9631 | 4.902 | 5.48 | 5.52 | 5.68 | 5.57 | 5.36
| Mean |  | 5.8767 | 5.6688 | 3.9369 | 5.532 | 5.9244 | 5.9233 | 5.9648 | 6.16 | 5.7017 | 5.8383 | 5.485 | 5.6067 | 5.695 | 5.7233 | 5.6427 | 5.8 | 5.7683 |  | 5.9933 | 5.1733 | 5.6783 | 6.8617 | 5.8683 | 4.973 | 5.4667 | 5.5633 | 5.7117 | 5.6 | 5.5617
| Median |  | 5.9 | 5.6425 | 3.9358 | 5.5239 | 5.8677 | 5.925 | 5.9695 | 6.18 | 5.715 | 5.825 | 5.485 | 5.61 | 5.7 | 5.72 | 5.6162 | 5.77 | 5.715 |  | 5.97 | 5.125 | 5.655 | 7.1 | 5.9516 | 4.9835 | 5.465 | 5.56 | 5.685 | 5.585 | 5.605
| Std Dev. |  | 0.045898 | 0.054064 | 0.036268 | 0.034108 | 0.16525 | 0.045461 | 0.056595 | 0.073756 | 0.083766 | 0.027869 | 0.040373 | 0.037771 | 0.041352 | 0.051251 | 0.17603 | 0.13038 | 0.12384 |  | 0.16896 | 0.13095 | 0.11822 | 0.78936 | 0.16265 | 0.079927 | 0.022509 | 0.066232 | 0.054191 | 0.046476 | 0.11409
| Rel.Std.Dev. |  | 0.781% | 0.9537% | 0.9212% | 0.6166% | 2.789% | 0.7675% | 0.9488% | 1.197% | 1.469% | 0.4773% | 0.7361% | 0.6737% | 0.7261% | 0.8955% | 3.12% | 2.248% | 2.147% |  | 2.819% | 2.531% | 2.082% | 11.5% | 2.772% | 1.607% | 0.4118% | 1.191% | 0.9488% | 0.8299% | 2.051%
| PDM3 |  | 2.404% | -1.218% | -31.4% | -3.603% | 3.236% | 3.217% | 3.94% | 7.341% | -0.6457% | 1.736% | -4.421% | -2.301% | -0.7619% | -0.2681% | -1.674% | 1.068% | 0.516% |  | 4.437% | -9.852% | -1.052% | 19.57% | 2.259% | -13.34% | -4.741% | -3.056% | -0.4714% | -2.417% | -3.085%
| Z-Score (Absolute) |  | 0.78 | 0.13 | 7.75 | 0.73 | 1 | 0.99 | 1.17 | 2.03 | 0.01 | 0.62 | 0.94 | 0.4 | 0.01 | 0.11 | 0.24 | 0.45 | 0.31 | NA | 1.3 | 2.31 | 0.09 | 5.12 | 0.75 | 3.19 | 1.02 | 0.59 | 0.06 | 0.43 | 0.6
| Analytical results for Ga in OREAS 233b (Certified Value 9.99 ppm) | [32 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 8.9 | 10.3 | 9.5924 | 8.7862 | 10.1 | 11.1 | 10.5 | 9.92 | 10 | 10 | 9.72 | 9.74 | 9.08 | 9.99 | 6 | 10.3 | 9.2 | 11.6 | 12.15 | 9 | 9.73 | 10.9 | 9.7 | 7.94 | 8.74 | [6 empty] | 1
| 1 | 2 | 9.3 | 10.4 | 9.6767 | 8.8008 | 9.94 | 11.7 | 10.3 | 10.22 | 10 | 11 | 9.78 | 10.05 | 9.41 | 9.98 | 7 | 9.5 | 9.7 | 11.6 | 11.98 | 9 | 9.65 | 11.7 | 10.2 | 8.06 | 8.83 | [6 empty] | 32
| 1 | 3 | 9.1 | 10.5 | 9.6933 | 9.376 | 9.92 | 10.6 | 10.6 | 10.02 | 10 | 10 | 9.96 | 10.3 | 9.54 | 9.79 | 6 | 10.4 | 9.3 | 11.5 | 12.64 | 9 | 9.69 | 11.4 | 10.5 | 8.01 | 8.88 | [6 empty] | 16
| 1 | 4 | 9.2 | 10.6 | 9.5261 | 8.9488 | 9.46 | 11.7 | 10.6 | 9.82 | 10 | 9 | 9.64 | 9.9 | 9.57 | 10.15 | 7 | 9.8 | 8.9 | 11.5 | 12.63 | 9 | 9.82 | 11.5 | 10.3 | 7.92 | 8.64 | [6 empty] | 9.9931
| 1 | 5 | 9.3 | 10.5 | 9.5935 | 8.7479 | 9.91 | 11 | 10.7 | 9.63 | 10 | 9 | 9.92 | 10.1 | 9.51 | 10.5 | 6 | 10.3 | 9.1 | 11.6 | 12.86 | 8 | 9.71 | 11.1 | 10.1 | 7.86 | 8.76 | [6 empty] | 93
| 1 | 6 | 9.2 | 10.4 | 9.6114 | 8.9403 | 10.2 | 11.8 | 10.6 | 9.99 | 10 | 10 | 9.78 | 10.2 | 9.89 | 10.35 | 7 | 9.7 | 9.1 | 11.2 | 13.23 | 8 | 9.58 | 11.1 | 9.8 | 7.91 | 8.58
| Mean |  | 9.1667 | 10.45 | 9.6156 | 8.9333 | 9.9217 | 11.3167 | 10.55 | 9.9333 | 10 | 9.8333 | 9.8 | 10.0483 | 9.5 | 10.1267 | 6.5 | 10 | 9.2167 | 11.5 | 12.5817 | 8.6667 | 9.6967 | 11.2833 | 10.1 | 7.95 | 8.7383
| Median |  | 9.2 | 10.45 | 9.6025 | 8.8705 | 9.93 | 11.4 | 10.6 | 9.955 | 10 | 10 | 9.78 | 10.075 | 9.525 | 10.07 | 6.5 | 10.05 | 9.15 | 11.55 | 12.635 | 9 | 9.7 | 11.25 | 10.15 | 7.93 | 8.75
| Std Dev. |  | 0.15055 | 0.10488 | 0.061395 | 0.23231 | 0.25412 | 0.48751 | 0.13784 | 0.19896 | 0 | 0.75277 | 0.12066 | 0.20302 | 0.26199 | 0.26204 | 0.54772 | 0.37947 | 0.27142 | 0.15492 | 0.45867 | 0.5164 | 0.080416 | 0.29944 | 0.30332 | 0.072664 | 0.11286
| Rel.Std.Dev. |  | 1.642% | 1.004% | 0.6385% | 2.6% | 2.561% | 4.308% | 1.307% | 2.003% | 0% | 7.655% | 1.231% | 2.02% | 2.758% | 2.588% | 8.427% | 3.795% | 2.945% | 1.347% | 3.646% | 5.958% | 0.8293% | 2.654% | 3.003% | 0.914% | 1.292%
| PDM3 |  | -8.27% | 4.572% | -3.778% | -10.61% | -0.7147% | 13.24% | 5.573% | -0.598% | 0.06916% | -1.599% | -1.932% | 0.5528% | -4.934% | 1.337% | -34.96% | 0.06916% | -7.77% | 15.08% | 25.9% | -13.27% | -2.966% | 12.91% | 1.07% | -20.45% | -12.56%
| Z-Score (Absolute) |  | 1 | 0.67 | 0.41 | 1.31 | 0.02 | 1.81 | 0.8 | 0 | NA | NA | 0.17 | 0.15 | 0.57 | 0.25 | NA | 0.09 | 0.94 | 2.04 | 3.46 | NA | 0.31 | 1.76 | 0.22 | 2.59 | 1.56
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 233b (Certified Value 1.8 ppm) | [32 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.27 |  |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.07: AR*MS |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g |  |  | 2
| 1 | 1 | 1.01: 1.918; 1.02: 1.581; 1.03: 1.901; 1.04: 2; 1.05: 1.9; 1.07: 1.6 |  |  | 1
| 1 | 2 | 1.01: 1.924; 1.02: 1.591; 1.03: 1.9372; 1.04: 2; 1.05: 1.86; 1.07: 1.6 |  |  | 33
| 1 | 3 | 1.01: 1.94; 1.02: 1.454; 1.03: 1.9; 1.04: 2; 1.05: 1.88; 1.07: 1.5 |  |  | 16
| 1 | 4 | 1.01: 1.941; 1.02: 1.423; 1.03: 1.8391; 1.04: 1.9; 1.05: 1.88; 1.07: 1.5 |  |  | 1.7984
| 1 | 5 | 1.01: 1.954; 1.02: 1.498; 1.03: 1.9202; 1.04: 2; 1.05: 1.77; 1.07: 1.5 |  |  | 94
| 1 | 6 | 1.01: 1.941; 1.02: 1.574; 1.03: 1.9521; 1.04: 2.1; 1.05: 1.94; 1.07: 1.5
| Mean |  | 1.01: 1.9363; 1.02: 1.5202; 1.03: 1.9083; 1.04: 2; 1.05: 1.8717; 1.07: 1.5333
| Median |  | 1.01: 1.9405; 1.02: 1.536; 1.03: 1.9106; 1.04: 2; 1.05: 1.88; 1.07: 1.5
| Std Dev. |  | 1.01: 0.013095; 1.02: 0.072009; 1.03: 0.039489; 1.04: 0.063246; 1.05: 0.056716; 1.07: 0.05164
| Rel.Std.Dev. |  | 1.01: 0.6763%; 1.02: 4.737%; 1.03: 2.069%; 1.04: 3.162%; 1.05: 3.03%; 1.07: 3.368%
| PDM3 |  | 1.01: 7.673%; 1.02: -15.47%; 1.03: 6.112%; 1.04: 11.21%; 1.05: 4.077%; 1.07: -14.74%
| Z-Score (Absolute) |  | 1.01: 0.4; 1.02: 3.19; 1.03: 0.16; 1.04: 0.95; 1.05: 0.16; 1.07: 3.08
| Analytical results for Ge in OREAS 233b (Certified Value 0.11 ppm) | [32 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 2
| 1 | 1 | < 0.1 | 0.08 | 0.2234 | < 0.2 | 0.1 | 0.14 | 2.1 | 0.12 | 0.14 | 0.09 | 0.12 | < 10 | < 0.1 | 0.15 | 0.1 | 0.2 | 0.1 | 0.1 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.08 | 0.22366 | < 0.2 | 0.1 | 0.12 | 2.4 | 0.11 | 0.14 | 0.1 | 0.13 | < 10 | < 0.1 | 0.17 | 0.11 | 0.2 | 0.1 | 0.1 | [13 empty] | 34
| 1 | 3 | < 0.1 | 0.09 | 0.23845 | < 0.2 | < 0.1 | 0.1 | 1.9 | 0.12 | 0.14 | 0.12 | 0.11 | < 10 | < 0.1 | 0.13 | 0.12 | 0.2 | 0.1 | 0.09 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.08 | 0.232 | < 0.2 | 0.1 | 0.12 | 2 | 0.11 | 0.14 | 0.1 | 0.12 | < 10 | < 0.1 | 0.14 | 0.07 | 0.2 | 0.1 | 0.11 | [13 empty] | 0.11367
| 1 | 5 | < 0.1 | 0.08 | 0.23174 | < 0.2 | < 0.1 | 0.11 | 2 | 0.11 | 0.15 | 0.1 | 0.12 | < 10 | 0.1 | 0.16 | 0.12 | 0.2 | 0.09 | 0.09 | [13 empty] | 95
| 1 | 6 | < 0.1 | 0.08 | 0.2213 | < 0.2 | < 0.1 | 0.11 | 2 | 0.12 | 0.14 | 0.1 | 0.12 | < 10 | < 0.1 | 0.15 | 0.16 | 0.2 | 0.1 | 0.1
| Mean |  | 1.02: 0.081667; 1.03: 0.22842; 1.05: 0.1; 1.07: 0.11667; 1.08: 2.0667; 1.10: 0.115; 1.11: 0.14167; 1.12: 0.10167; 1.13: 0.12; 1.15: 0.1; 1.16: 0.15; 1.17: 0.11333; 1.18: 0.2; 1.19: 0.098333; 1.21: 0.098333
| Median |  | 1.02: 0.08; 1.03: 0.2277; 1.05: 0.1; 1.07: 0.115; 1.08: 2; 1.10: 0.115; 1.11: 0.14; 1.12: 0.1; 1.13: 0.12; 1.15: 0.1; 1.16: 0.15; 1.17: 0.115; 1.18: 0.2; 1.19: 0.1; 1.21: 0.1
| Std Dev. |  | 1.02: 0.0040825; 1.03: 0.0066776; 1.05: 1.6997e-17; 1.07: 0.013663; 1.08: 0.17512; 1.10: 0.0054772; 1.11: 0.0040825; 1.12: 0.0098319; 1.13: 0.0063246; 1.16: 0.014142; 1.17: 0.029439; 1.18: 3.0405e-17; 1.19: 0.0040825; 1.21: 0.0075277
| Rel.Std.Dev. |  | 1.02: 4.999%; 1.03: 2.923%; 1.05: 1.7e-14%; 1.07: 11.71%; 1.08: 8.474%; 1.10: 4.763%; 1.11: 2.882%; 1.12: 9.671%; 1.13: 5.27%; 1.16: 9.428%; 1.17: 25.98%; 1.18: 1.52e-14%; 1.19: 4.152%; 1.21: 7.655%
| PDM3 |  | 1.02: -28.15%; 1.03: 101%; 1.05: -12.02%; 1.07: 2.639%; 1.08: 1718%; 1.10: 1.173%; 1.11: 24.63%; 1.12: -10.56%; 1.13: 5.572%; 1.15: -12.02%; 1.16: 31.96%; 1.17: -0.2933%; 1.18: 75.95%; 1.19: -13.49%; 1.21: -13.49%
| Z-Score (Absolute) |  | 2.73 | 1.38 | 4.87 | 0.6 | NA | 0.11 | NA | 0.04 | 1.17 | 0.53 | 0.25 | 208.08 | NA | 1.53 | 0.04 | NA | 0.67 | 0.67
Results from laboratories 1.07, 1.13, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 233b (Certified Value 0.41 ppm) | [32 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 2
| 1 | 1 | 0.31 | 0.44 | 0.36706 | 0.4 | 0.35 | 0.43 | 0.56 | 0.3 | 0.42 | 0.43 | 0.39 | 0.38 | 0.47 | < 0.5 | 0.43 | 0.72 | 0.41 | 0.47 | 0.39 | [12 empty] | 1
| 1 | 2 | 0.31 | 0.44 | 0.38388 | 0.4 | 0.38 | 0.41 | 0.55 | 0.3 | 0.42 | 0.47 | 0.44 | 0.4 | 0.44 | < 0.5 | 0.43 | 0.71 | 0.43 | 0.48 | 0.38 | [12 empty] | 18
| 1 | 3 | 0.31 | 0.44 | 0.39044 | 0.3 | 0.35 | 0.42 | 0.58 | 0.3 | 0.48 | 0.46 | 0.43 | 0.4 | 0.44 | < 0.5 | 0.41 | 0.73 | 0.42 | 0.46 | 0.37 | [12 empty] | 16
| 1 | 4 | 0.36 | 0.44 | 0.3602 | 0.4 | 0.33 | 0.43 | 0.55 | 0.3 | 0.45 | 0.44 | 0.44 | 0.41 | 0.46 | < 0.5 | 0.44 | 0.67 | 0.43 | 0.48 | 0.37 | [12 empty] | 0.41292
| 1 | 5 | 0.34 | 0.44 | 0.37436 | 0.4 | 0.32 | 0.43 | 0.57 | 0.3 | 0.42 | 0.46 | 0.45 | 0.37 | 0.47 | < 0.5 | 0.43 | 0.68 | 0.46 | 0.45 | 0.39 | [12 empty] | 96
| 1 | 6 | 0.33 | 0.42 | 0.40904 | 0.4 | 0.33 | 0.4 | 0.56 | 0.3 | 0.45 | 0.47 | 0.41 | 0.39 | 0.46 | < 0.5 | 0.42 | 0.62 | 0.43 | 0.47 | 0.37
| Mean |  | 0.32667 | 0.43667 | 0.38083 | 0.38333 | 0.34333 | 0.42 | 0.56167 | 0.3 | 0.44 | 0.455 | 0.42667 | 0.39167 | 0.45667 |  | 0.42667 | 0.68833 | 0.43 | 0.46833 | 0.37833
| Median |  | 0.32 | 0.44 | 0.37912 | 0.4 | 0.34 | 0.425 | 0.56 | 0.3 | 0.435 | 0.46 | 0.435 | 0.395 | 0.46 |  | 0.43 | 0.695 | 0.43 | 0.47 | 0.375
| Std Dev. |  | 0.020656 | 0.008165 | 0.017632 | 0.040825 | 0.021602 | 0.012649 | 0.01169 | 0 | 0.024495 | 0.016432 | 0.022509 | 0.01472 | 0.013663 |  | 0.010328 | 0.040702 | 0.016733 | 0.01169 | 0.0098319
| Rel.Std.Dev. |  | 6.323% | 1.87% | 4.63% | 10.65% | 6.292% | 3.012% | 2.081% | 0% | 5.567% | 3.611% | 5.276% | 3.758% | 2.992% |  | 2.421% | 5.913% | 3.891% | 2.496% | 2.599%
| PDM3 |  | -20.89% | 5.752% | -7.771% | -7.164% | -16.85% | 1.716% | 36.02% | -27.35% | 6.559% | 10.19% | 3.33% | -5.146% | 10.6% |  | 3.33% | 66.7% | 4.137% | 13.42% | -8.375%
| Z-Score (Absolute) |  | 1.93 | 0.19 | 0.88 | NA | 1.61 | 0.13 | 2.6 | NA | 0.26 | 0.55 | 0 | 0.67 | 0.58 | 3.4 | 0 | 5.04 | 0.06 | 0.8 | 0.93
Results from laboratories 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 233b (Indicative Value 0.045 ppm) | [32 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | < 0.01 | 0.02 | < 1 | 0.06 | < 1 | < 0.1 | < 5 | 0.656 | 0.02 | 0.02 | < 0.01 | 0.04 | < 1 | 0.08 | 0.012 | 0.07 | 0.27 | 0.06 | < 0.08 | 0.03 | 0.11 | 0.26 | < 0.01 | 0.02 | 0.35 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.02 | < 1 | 0.07 | < 1 | < 0.1 | < 5 | 0.715 | 0.01 | 0.03 | < 0.01 | 0.04 | < 1 | 0.11 | 0.013 | 0.05 | 0.28 | 0.05 | < 0.08 | 0.04 | 0.098 | 0.27 | < 0.01 | 0.023 | 0.35 | [6 empty] | 5
| 1 | 3 | < 0.01 | 0.02 | < 1 | 0.06 | < 1 | < 0.1 | < 5 | 0.614 | 0.01 | 0.02 | < 0.01 | 0.03 | < 1 | 0.1 | 0.01 | 0.08 | 0.15 | 0.05 | < 0.08 | 0.03 | 0.106 | 0.26 | < 0.01 | 0.022 | 0.34 | [6 empty] | 16
| 1 | 4 | < 0.01 | 0.02 | < 1 | 0.09 | < 1 | < 0.1 | < 5 | 0.62 | 0.01 | 0.03 | < 0.01 | 0.02 | < 1 | 0.09 | 0.014 | 0.08 | 0.48 | 0.05 | < 0.08 | 0.03 | 0.094 | 0.27 | < 0.01 | 0.02 | 0.33 | [6 empty] | 0.044556
| 1 | 5 | < 0.01 | 0.02 | < 1 | 0.08 | < 1 | < 0.1 | < 5 | 0.603 | 0.02 | 0.03 | < 0.01 | 0.05 | < 1 | 0.09 | 0.012 | 0.05 | < 0.1 | 0.05 | < 0.08 | 0.03 | 0.095 | 0.26 | < 0.01 | 0.021 | 0.36 | [6 empty] | 11
| 1 | 6 | < 0.01 | 0.02 | < 1 | 0.05 | < 1 | < 0.1 | < 5 | 0.602 | 0.02 | 0.02 | < 0.01 | 0.03 | < 1 | 0.08 | 0.01 | 0.05 | < 0.1 | 0.05 | < 0.08 | 0.03 | 0.094 | 0.29 | < 0.01 | 0.024 | 0.35
| Mean |  |  | 0.02 |  | 0.068333 |  |  |  | 0.635 | 0.015 | 0.025 |  | 0.035 |  | 0.091667 | 0.011833 | 0.063333 | 0.295 | 0.051667 |  | 0.031667 | 0.0995 | 0.26833 |  | 0.021667 | 0.34667
| Median |  |  | 0.02 |  | 0.065 |  |  |  | 0.617 | 0.015 | 0.025 |  | 0.035 |  | 0.09 | 0.012 | 0.06 | 0.275 | 0.05 |  | 0.03 | 0.0965 | 0.265 |  | 0.0215 | 0.35
| Std Dev. |  |  | 0 |  | 0.01472 |  |  |  | 0.043863 | 0.0054772 | 0.0054772 |  | 0.010488 |  | 0.01169 | 0.0016021 | 0.015055 | 0.13675 | 0.0040825 |  | 0.0040825 | 0.0068629 | 0.01169 |  | 0.001633 | 0.010328
| Rel.Std.Dev. |  |  | 0% |  | 21.54% |  |  |  | 6.908% | 36.51% | 21.91% |  | 29.97% |  | 12.75% | 13.54% | 23.77% | 46.36% | 7.902% |  | 12.89% | 6.897% | 4.357% |  | 7.537% | 2.979%
| PDM3 |  |  | -55.11% |  | 53.37% |  |  |  | 1325% | -66.33% | -43.89% |  | -21.45% |  | 105.7% | -73.44% | 42.14% | 562.1% | 15.96% |  | -28.93% | 123.3% | 502.2% |  | -51.37% | 678.1%
| Z-Score (Absolute) |  | 0.79 | 0.53 | 7.56 | 0.28 | 7.56 | 0.03 | 41.27 | 9.83 | 0.62 | 0.45 | 0.79 | 0.28 | 7.56 | 0.67 | 0.67 | 0.2 | 2.73 | 0 | 0.2 | 0.34 | 0.81 | 3.65 | 0.79 | 0.51 | 4.97
| Analytical results for Ho in OREAS 233b (Certified Value 0.42 ppm) | [32 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.11: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.05: 1.04; 1.07: 1.07; 1.08: 1.08; 1.11: 1.27 | [6 empty] | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.05: AR*MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.11: AR*MS | [6 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.05: 0.5g; 1.07: 0.5g; 1.08: 0.5g; 1.11: 25g | [6 empty] | 2
| 1 | 1 | 1.01: 0.411; 1.02: 0.417; 1.05: 0.37758; 1.07: 0.5; 1.08: 0.54; 1.11: 0.35 | [6 empty] | 1
| 1 | 2 | 1.01: 0.405; 1.02: 0.413; 1.05: 0.38808; 1.07: 0.5; 1.08: 0.52; 1.11: 0.35 | [6 empty] | 19
| 1 | 3 | 1.01: 0.416; 1.02: 0.379; 1.05: 0.38232; 1.07: 0.5; 1.08: 0.52; 1.11: 0.3 | [6 empty] | 16
| 1 | 4 | 1.01: 0.42; 1.02: 0.381; 1.05: 0.35482; 1.07: 0.4; 1.08: 0.53; 1.11: 0.3 | [6 empty] | 0.42261
| 1 | 5 | 1.01: 0.421; 1.02: 0.395; 1.05: 0.37816; 1.07: 0.5; 1.08: 0.51; 1.11: 0.35 | [6 empty] | 97
| 1 | 6 | 1.01: 0.412; 1.02: 0.416; 1.05: 0.39715; 1.07: 0.5; 1.08: 0.53; 1.11: 0.35
| Mean |  | 1.01: 0.41417; 1.02: 0.40017; 1.05: 0.37968; 1.07: 0.48333; 1.08: 0.525; 1.11: 0.33333
| Median |  | 1.01: 0.414; 1.02: 0.404; 1.05: 0.38024; 1.07: 0.5; 1.08: 0.525; 1.11: 0.35
| Std Dev. |  | 1.01: 0.006047; 1.02: 0.017555; 1.05: 0.014195; 1.07: 0.040825; 1.08: 0.010488; 1.11: 0.02582
| Rel.Std.Dev. |  | 1.01: 1.46%; 1.02: 4.387%; 1.05: 3.739%; 1.07: 8.447%; 1.08: 1.998%; 1.11: 7.746%
| PDM3 |  | 1.01: -1.999%; 1.02: -5.312%; 1.05: -10.16%; 1.07: 14.37%; 1.08: 24.23%; 1.11: -21.13%
| Z-Score (Absolute) |  | 1.01: 0.09; 1.02: 0.09; 1.05: 0.37; 1.07: 1.01; 1.08: 1.57; 1.11: 0.98
| Analytical results for In in OREAS 233b (Certified Value 0.03 ppm) | [32 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.03 | 0.026433 | 0.02 | < 0.05 | 0.04 | < 5 | 0.04 | 0.028 | 0.032 | 0.026 | 0.031 | < 10 | < 0.05 | 2.66 | < 0.1 | 0.03376 | 0.028 | 0.029 | 0.03 | 0.03 | 0.026 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.03 | 0.028752 | 0.03 | < 0.05 | 0.03 | < 5 | 0.04 | 0.032 | 0.031 | 0.03 | 0.029 | < 10 | < 0.05 | 1.23 | < 0.1 | 0.03107 | 0.031 | 0.032 | 0.03 | 0.031 | 0.027 | [9 empty] | 20
| 1 | 3 | 0.03 | 0.03 | 0.029084 | 0.03 | < 0.05 | 0.04 | < 5 | 0.03 | 0.032 | 0.03 | 0.031 | 0.028 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.03139 | 0.03 | 0.034 | 0.03 | 0.029 | 0.023 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.03 | 0.026897 | 0.03 | < 0.05 | 0.03 | < 5 | 0.03 | 0.032 | 0.03 | 0.028 | 0.03 | < 10 | < 0.05 | 0.47 | < 0.1 | 0.03168 | 0.029 | 0.029 | 0.03 | 0.03 | 0.026 | [9 empty] | 0.030076
| 1 | 5 | 0.03 | 0.03 | 0.028089 | 0.03 | < 0.05 | 0.03 | < 5 | 0.03 | 0.034 | 0.032 | 0.028 | 0.03 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.02861 | 0.032 | 0.031 | 0.03 | 0.028 | 0.028 | [9 empty] | 98
| 1 | 6 | 0.03 | 0.03 | 0.030309 | 0.02 | < 0.05 | 0.03 | < 5 | 0.04 | 0.032 | 0.029 | 0.032 | 0.033 | < 10 | < 0.05 | 0.44 | < 0.1 | 0.0288 | 0.028 | 0.034 | 0.03 | 0.028 | 0.025
| Mean |  | 0.028333 | 0.03 | 0.02826 | 0.026667 |  | 0.033333 |  | 0.035 | 0.031667 | 0.030667 | 0.029167 | 0.030167 |  |  | 1.2 |  | 0.030885 | 0.029667 | 0.0315 | 0.03 | 0.029333 | 0.025833
| Median |  | 0.03 | 0.03 | 0.02842 | 0.03 |  | 0.03 |  | 0.035 | 0.032 | 0.0305 | 0.029 | 0.03 |  |  | 0.85 |  | 0.03123 | 0.0295 | 0.0315 | 0.03 | 0.0295 | 0.026
| Std Dev. |  | 0.0040825 | 0 | 0.0014382 | 0.005164 |  | 0.005164 |  | 0.0054772 | 0.0019664 | 0.0012111 | 0.0022286 | 0.0017224 |  |  | 1.0397 |  | 0.0019344 | 0.001633 | 0.0022583 | 0 | 0.0012111 | 0.0017224
| Rel.Std.Dev. |  | 14.41% | 0% | 5.089% | 19.36% |  | 15.49% |  | 15.65% | 6.21% | 3.949% | 7.641% | 5.71% |  |  | 86.64% |  | 6.263% | 5.504% | 7.169% | 0% | 4.129% | 6.667%
| PDM3 |  | -5.793% | -0.2519% | -6.036% | -11.34% |  | 10.83% |  | 16.37% | 5.29% | 1.965% | -3.023% | 0.3023% |  |  | 3890% |  | 2.691% | -1.36% | 4.736% | -0.2519% | -2.469% | -14.11%
| Z-Score (Absolute) |  | 0.66 | 0.03 | 0.69 | 1.29 | 1.92 | 1.23 | 932.3 | 1.86 | 0.6 | 0.22 | 0.35 | 0.03 | 1875.95 | 1.92 | 296.9 | 7.52 | 0.3 | 0.16 | 0.53 | 0.03 | 0.28 | 1.6
| Analytical results for Ir in OREAS 233b (Indicative Value < 0.003 ppm) | [32 empty] | Indicative
Ir | Round | Replicate | 1.01: Lab; 1.02: Lab | [6 empty] | 1
| No | No | 1.01: 1.27; 1.02: 1.28 | [6 empty] | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS | [6 empty] | 3
|  |  | 1.01: 25g; 1.02: 0.5g | [6 empty] | 3
| 1 | 1 | 1.01: N.A.; 1.02: < 0.003 | [6 empty] | 1
| 1 | 2 | 1.01: N.A.; 1.02: < 0.003 | [6 empty] | 6
| 1 | 3 | 1.01: N.A.; 1.02: < 0.003 | [6 empty] | 16
| 1 | 4 | 1.01: N.A.; 1.02: < 0.003 | [6 empty] | < 0.003
| 1 | 5 | 1.01: N.A.; 1.02: < 0.003 | [6 empty] | 12
| 1 | 6 | 1.01: N.A.; 1.02: < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.02
| Analytical results for K in OREAS 233b (Certified Value 0.143 wt.%) | [32 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.14 | 0.131 | 0.19177 | 0.15247 | 0.14146 | 0.13 | 0.1552 | 0.163 | 0.1456 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14141 | 0.14 | 0.12 | 0.14623 | 0.144 | 0.14 | 0.15 | 0.13703 | 0.13558 | 0.126 | 0.13 | 0.14 | 0.17 | 0.14 | 0.14 |  |  | 1
| 1 | 2 | 0.15 | 0.13 | 0.18624 | 0.15841 | 0.13863 | 0.12 | 0.1511 | 0.162 | 0.1468 | 0.14 | 0.15 | 0.14 | 0.15 | 0.14 | 0.1474 | 0.14 | 0.13 | 0.15014 | 0.147 | 0.15 | 0.15 | 0.13414 | 0.13784 | 0.1285 | 0.13 | 0.15 | 0.17 | 0.15 | 0.14 |  |  | 0
| 1 | 3 | 0.15 | 0.133 | 0.18642 | 0.15911 | 0.15201 | 0.12 | 0.1544 | 0.167 | 0.1451 | 0.14 | 0.15 | 0.14 | 0.14 | 0.14 | 0.15121 | 0.14 | 0.13 | 0.154 | 0.146 | 0.14 | 0.14 | 0.13918 | 0.14339 | 0.123 | 0.13 | 0.14 | 0.16 | 0.14 | 0.14 |  |  | 16
| 1 | 4 | 0.15 | 0.131 | 0.18363 | 0.15034 | 0.13995 | 0.12 | 0.1511 | 0.182 | 0.1428 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14533 | 0.14 | 0.13 | 0.15208 | 0.142 | 0.14 | 0.15 | 0.1381 | 0.14587 | 0.1265 | 0.13 | 0.15 | 0.16 | 0.14 | 0.14 |  |  | 0.14303
| 1 | 5 | 0.15 | 0.131 | 0.18307 | 0.15772 | 0.13997 | 0.12 | 0.1527 | 0.171 | 0.144 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14061 | 0.14 | 0.13 | 0.15312 | 0.149 | 0.14 | 0.14 | 0.13679 | 0.14238 | 0.123 | 0.13 | 0.14 | 0.17 | 0.14 | 0.14 |  |  | 99
| 1 | 6 | 0.15 | 0.13 | 0.19056 | 0.15776 | 0.14123 | 0.12 | 0.1536 | 0.165 | 0.1484 | 0.15 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14112 | 0.14 | 0.13 | 0.14999 | 0.141 | 0.15 | 0.14 | 0.13919 | 0.14183 | 0.1205 | 0.13 | 0.14 | 0.17 | 0.14 | 0.14
| Mean |  | 0.14833 | 0.131 | 0.18695 | 0.15597 | 0.14221 | 0.12167 | 0.15302 | 0.16833 | 0.14545 | 0.14667 | 0.14333 | 0.14 | 0.14333 | 0.14 | 0.14451 | 0.14 | 0.12833 | 0.15093 | 0.14483 | 0.14333 | 0.145 | 0.1374 | 0.14115 | 0.12458 | 0.13 | 0.14333 | 0.16667 | 0.14167 | 0.14
| Median |  | 0.15 | 0.131 | 0.18633 | 0.15774 | 0.1406 | 0.12 | 0.15315 | 0.166 | 0.14535 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14337 | 0.14 | 0.13 | 0.15111 | 0.145 | 0.14 | 0.145 | 0.13756 | 0.14211 | 0.1245 | 0.13 | 0.14 | 0.17 | 0.14 | 0.14
| Std Dev. |  | 0.0040825 | 0.0010954 | 0.0035552 | 0.0036339 | 0.0049106 | 0.0040825 | 0.0017011 | 0.0074207 | 0.0019897 | 0.005164 | 0.005164 | 0 | 0.005164 | 0 | 0.0042475 | 0 | 0.0040825 | 0.0027978 | 0.0030605 | 0.005164 | 0.0054772 | 0.001898 | 0.003775 | 0.0029226 | 0 | 0.005164 | 0.005164 | 0.0040825 | 0
| Rel.Std.Dev. |  | 2.752% | 0.8362% | 1.902% | 2.33% | 3.453% | 3.355% | 1.112% | 4.408% | 1.368% | 3.521% | 3.603% | 0% | 3.603% | 0% | 2.939% | 0% | 3.181% | 1.854% | 2.113% | 3.603% | 3.777% | 1.381% | 2.674% | 2.346% | 0% | 3.603% | 3.098% | 2.882% | 0%
| PDM3 |  | 3.711% | -8.408% | 30.71% | 9.049% | -0.5707% | -14.93% | 6.985% | 17.69% | 1.695% | 2.545% | 0.2147% | -2.116% | 0.2147% | -2.116% | 1.04% | -2.116% | -10.27% | 5.525% | 1.263% | 0.2147% | 1.38% | -3.931% | -1.313% | -12.89% | -9.108% | 0.2147% | 16.53% | -0.9506% | -2.116%
| Z-Score (Absolute) |  | 1.01 | 2.49 | 8.82 | 2.56 | 0.23 | 4.38 | 1.96 | 5.06 | 0.43 | 0.67 | 0 | 0.67 | 0 | 0.67 | 0.24 | 0.67 | 3.03 | 1.54 | 0.3 | 0 | 0.34 | 1.2 | 0.44 | 3.79 | 2.7 | 0 | 4.72 | 0.34 | 0.67
| Analytical results for La in OREAS 233b (Certified Value 4.09 ppm) | [32 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 3.8 | 4.134 | 4.355 | 3.5756 | 4.6655 | 4.3 | 1 | 4.64 | 4.158 | < 5 | 4 | 3.9 | 4.2 | 3.9 | 4.1 | 4 | 3.96 | 3.8 | 4 | 4.5 | 4.2 | 4.4 | 4.1 | 4.2 | 4.3 | [6 empty] | 1
| 1 | 2 | 3.8 | 4.098 | 3.764 | 3.7241 | 4.6237 | 4.2 | < 1 | 4.64 | 4.123 | < 5 | 5 | 3.9 | 4.2 | 4 | 4.1 | 4 | 4.14 | 3.8 | 4.1 | 4 | 4.2 | 4.5 | 4.2 | 4.3 | 4.3 | [6 empty] | 21
| 1 | 3 | 3.9 | 4.163 | 3.529 | 3.6894 | 4.9033 | 4 | 1 | 4.81 | 4.105 | < 5 | 4 | 3.9 | 4.1 | 4 | 4 | 4 | 4.15 | 3.7 | 4.2 | 4 | 4.1 | 4.4 | 4.1 | 4.3 | 4.3 | [6 empty] | 16
| 1 | 4 | 3.8 | 4.151 | 3.51 | 3.7531 | 4.6391 | 4 | < 1 | 4.99 | 3.947 | < 5 | 4 | 3.8 | 4.2 | 4.1 | 4.1 | 4 | 3.91 | 3.8 | 4.1 | 4 | 4.2 | 4.4 | 4.2 | 4.3 | 4.2 | [6 empty] | 4.0936
| 1 | 5 | 3.9 | 4.18 | 3.629 | 3.5427 | 4.4882 | 4.2 | 1 | 4.91 | 4.113 | < 5 | 4 | 3.9 | 4.2 | 4.1 | 4.3 | 4 | 4.03 | 3.7 | 4.4 | 4 | 4.1 | 4.6 | 4.1 | 4.2 | 4.2 | [6 empty] | 100
| 1 | 6 | 3.9 | 4.131 | 3.736 | 3.7128 | 4.7686 | 4.2 | 1 | 4.67 | 4.126 | < 5 | 4 | 3.8 | 4.3 | 4.2 | 4.2 | 4 | 4.05 | 3.7 | 4.1 | 4 | 4.1 | 4.4 | 3.9 | 4.2 | 4.1
| Mean |  | 3.85 | 4.1428 | 3.7538 | 3.6663 | 4.6814 | 4.15 | 1 | 4.7767 | 4.0953 |  | 4.1667 | 3.8667 | 4.2 | 4.05 | 4.1333 | 4 | 4.04 | 3.75 | 4.15 | 4.0833 | 4.15 | 4.45 | 4.1 | 4.25 | 4.2333
| Median |  | 3.85 | 4.1425 | 3.6825 | 3.7011 | 4.6523 | 4.2 | 1 | 4.74 | 4.118 |  | 4 | 3.9 | 4.2 | 4.05 | 4.1 | 4 | 4.04 | 3.75 | 4.1 | 4 | 4.15 | 4.4 | 4.1 | 4.25 | 4.25
| Std Dev. |  | 0.054772 | 0.028576 | 0.31222 | 0.086122 | 0.14112 | 0.12247 | 0 | 0.15042 | 0.074888 |  | 0.40825 | 0.05164 | 0.063246 | 0.10488 | 0.10328 | 0 | 0.095499 | 0.054772 | 0.13784 | 0.20412 | 0.054772 | 0.083666 | 0.10954 | 0.054772 | 0.08165
| Rel.Std.Dev. |  | 1.423% | 0.6898% | 8.317% | 2.349% | 3.014% | 2.951% | 0% | 3.149% | 1.829% |  | 9.798% | 1.336% | 1.506% | 2.59% | 2.499% | 0% | 2.364% | 1.461% | 3.321% | 4.999% | 1.32% | 1.88% | 2.672% | 1.289% | 1.929%
| PDM3 |  | -5.951% | 1.203% | -8.3% | -10.44% | 14.36% | 1.378% | -75.57% | 16.69% | 0.04239% |  | 1.785% | -5.544% | 2.599% | -1.065% | 0.9707% | -2.286% | -1.309% | -8.394% | 1.378% | -0.2508% | 1.378% | 8.706% | 0.1564% | 3.821% | 3.414%
| Z-Score (Absolute) |  | 1.26 | 0.22 | 1.75 | 2.19 | 2.94 | 0.25 | 16.52 | 3.42 | 0.02 | 8.09 | NA | 1.18 | 0.51 | 0.25 | 0.17 | NA | 0.3 | 1.77 | 0.25 | 0.08 | 0.25 | 1.77 | 0 | 0.76 | 0.67
Results from laboratories 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 233b (Certified Value 10.5 ppm) | [32 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 1 | 10.04 | 16.75 | 4.1262 | 10.508 | 10.2 | 6 | 11.9 | 10.5 | 11 | 15.3 | 10.9 | 10.8 | 11.4 | 10.7 | < 5 | 8.2625 | 10.3 | 10 | 9 | 10.7 | 9.9 | 12 | 11.7 | 10 | [6 empty] | 1
| 1 | 2 | < 1 | 9.99 | 16.87 | 4.1379 | 10.4793 | 9.9 | 6 | 11.5 | 10.5 | 11 | 14.7 | 10.9 | 10.8 | 11.4 | 10.6 | < 5 | 8.797 | 8.9 | 10 | 9 | 10.8 | 10.3 | 12 | 12 | 10.2 | [6 empty] | 0
| 1 | 3 | < 1 | 10.3 | 17.5 | 4.162 | 11.18 | 10 | 6 | 11.7 | 10.2 | 11 | 14.9 | 10.8 | 10.8 | 11.4 | 10.5 | < 5 | 8.9701 | 10.3 | 10 | 9 | 10.9 | 9.9 | 11 | 11.4 | 9.6 | [6 empty] | 16
| 1 | 4 | < 1 | 10.04 | 16.69 | 3.9709 | 10.4042 | 9.6 | 6 | 11.9 | 9.9 | 11 | 14.4 | 10.9 | 11 | 11.2 | 10.8 | < 5 | 11.3902 | 9.4 | 10 | 9 | 10.7 | 10.1 | 12 | 11.4 | 9.2 | [6 empty] | 10.4844
| 1 | 5 | < 1 | 10.13 | 16.71 | 4.1391 | 10.4583 | 10 | 6 | 11.7 | 10.4 | 11 | 14 | 11 | 11.8 | 11.6 | 11.1 | < 5 | 8.2397 | 10.1 | 10 | 9 | 10.9 | 10.2 | 12 | 11.3 | 9.9 | [6 empty] | 101
| 1 | 6 | 1 | 9.92 | 17.22 | 4.0254 | 10.5072 | 10.3 | 6 | 11.7 | 10.5 | 11 | 13.4 | 10.8 | 11.1 | 11.6 | 11 | < 5 | 9.5904 | 9.1 | 10 | 9 | 10.8 | 10 | 11 | 11 | 9.8
| Mean |  | 1 | 10.07 | 16.9567 | 4.0936 | 10.5895 | 10 | 6 | 11.7333 | 10.3333 | 11 | 14.45 | 10.8833 | 11.05 | 11.4333 | 10.7833 |  | 9.2083 | 9.6833 | 10 | 9 | 10.8 | 10.0667 | 11.6667 | 11.4667 | 9.7833
| Median |  | 1 | 10.04 | 16.81 | 4.132 | 10.4933 | 10 | 6 | 11.7 | 10.45 | 11 | 14.55 | 10.9 | 10.9 | 11.4 | 10.75 |  | 8.8835 | 9.75 | 10 | 9 | 10.8 | 10.05 | 12 | 11.4 | 9.85
| Std Dev. |  | 0 | 0.13206 | 0.33068 | 0.076786 | 0.29182 | 0.24495 | 0 | 0.15055 | 0.24221 | 0 | 0.6775 | 0.075277 | 0.38859 | 0.15055 | 0.23166 |  | 1.1803 | 0.62743 | 0 | 0 | 0.089443 | 0.1633 | 0.5164 | 0.34448 | 0.34881
| Rel.Std.Dev. |  | 0% | 1.311% | 1.95% | 1.876% | 2.756% | 2.449% | 0% | 1.283% | 2.344% | 0% | 4.689% | 0.6917% | 3.517% | 1.317% | 2.148% |  | 12.82% | 6.479% | 0% | 0% | 0.8282% | 1.622% | 4.426% | 3.004% | 3.565%
| PDM3 |  | -90.46% | -3.953% | 61.73% | -60.96% | 1.002% | -4.621% | -42.77% | 11.91% | -1.441% | 4.917% | 37.82% | 3.805% | 5.394% | 9.05% | 2.851% |  | -12.17% | -7.641% | -4.621% | -14.16% | 3.01% | -3.985% | 11.28% | 9.368% | -6.687%
| Z-Score (Absolute) |  | NA | 0.43 | 5.33 | 5.43 | 0 | 0.49 | NA | 0.96 | 0.21 | NA | 3.23 | 0.25 | 0.39 | 0.71 | 0.16 | 6.77 | 1.16 | 0.76 | NA | NA | 0.18 | 0.44 | NA | 0.73 | 0.67
Results from laboratories 1.01, 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 233b (Certified Value 0.13 ppm) | [32 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab | [6 empty] | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.27; 1.08: 1.32 | [6 empty] | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.08: AR*OES/MS | [6 empty] | 2
|  |  | 1.01: 0.25g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g; 1.08: 0.25g | [6 empty] | 2
| 1 | 1 | 1.01: 0.12; 1.02: 0.146; 1.03: 0.139; 1.04: 0.34572; 1.05: 0.2; 1.07: 0.1; 1.08: 0.13 | [6 empty] | 1
| 1 | 2 | 1.01: 0.12; 1.02: 0.142; 1.03: 0.137; 1.04: 0.34869; 1.05: 0.2; 1.07: 0.1; 1.08: 0.13 | [6 empty] | 22
| 1 | 3 | 1.01: 0.12; 1.02: 0.147; 1.03: 0.127; 1.04: 0.33892; 1.05: 0.2; 1.07: < 0.1; 1.08: 0.13 | [6 empty] | 16
| 1 | 4 | 1.01: 0.13; 1.02: 0.15; 1.03: 0.127; 1.04: 0.33276; 1.05: 0.2; 1.07: 0.1; 1.08: 0.14 | [6 empty] | 0.12723
| 1 | 5 | 1.01: 0.13; 1.02: 0.149; 1.03: 0.134; 1.04: 0.32787; 1.05: 0.2; 1.07: < 0.1; 1.08: 0.13 | [6 empty] | 102
| 1 | 6 | 1.01: 0.12; 1.02: 0.147; 1.03: 0.142; 1.04: 0.33471; 1.05: 0.2; 1.07: < 0.1; 1.08: 0.13
| Mean |  | 1.01: 0.12333; 1.02: 0.14683; 1.03: 0.13433; 1.04: 0.33811; 1.05: 0.2; 1.07: 0.1; 1.08: 0.13167
| Median |  | 1.01: 0.12; 1.02: 0.147; 1.03: 0.1355; 1.04: 0.33682; 1.05: 0.2; 1.07: 0.1; 1.08: 0.13
| Std Dev. |  | 1.01: 0.005164; 1.02: 0.0027869; 1.03: 0.0062503; 1.04: 0.0079433; 1.05: 3.0405e-17; 1.07: 1.6997e-17; 1.08: 0.0040825
| Rel.Std.Dev. |  | 1.01: 4.187%; 1.02: 1.898%; 1.03: 4.653%; 1.04: 2.349%; 1.05: 1.52e-14%; 1.07: 1.7e-14%; 1.08: 3.101%
| PDM3 |  | 1.01: -3.065%; 1.02: 15.4%; 1.03: 5.58%; 1.04: 165.7%; 1.05: 57.19%; 1.07: -21.4%; 1.08: 3.484%
| Z-Score (Absolute) |  | 1.01: 0.59; 1.02: 0.67; 1.03: 0; 1.04: 10.99; 1.05: 3.54; 1.07: 3.2; 1.08: 0.14
| Analytical results for Mg in OREAS 233b (Certified Value 1.81 wt.%) | [32 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.85 | 1.778 | 1.8046 | 1.8968 | 1.8322 | 1.97 | 2.0263 | 1.99 | 1.84 | 1.89 | 1.75 | 1.8 | 1.78 | 1.75 | 1.7492 | 1.86 | 1.68 | 1.8641 | 1.92 | 1.75 | 1.83 | 2.5053 | 1.5548 | 1.644 | 1.76 | 1.77 | 1.8 | 1.77 | 1.87 |  |  | 1
| 1 | 2 | 1.87 | 1.737 | 1.7591 | 1.9053 | 1.8266 | 1.89 | 1.9901 | 1.94 | 1.82 | 1.86 | 1.73 | 1.79 | 1.82 | 1.75 | 1.8267 | 1.86 | 1.76 | 1.8991 | 1.87 | 1.85 | 1.81 | 2.7425 | 1.5543 | 1.6555 | 1.74 | 1.83 | 1.77 | 1.8 | 1.88 |  |  | 0
| 1 | 3 | 1.87 | 1.788 | 1.7688 | 1.9128 | 1.8503 | 1.91 | 2.0142 | 1.99 | 1.82 | 1.83 | 1.73 | 1.78 | 1.8 | 1.76 | 1.8795 | 1.83 | 1.73 | 1.9343 | 1.83 | 1.79 | 1.8 | 2.0629 | 1.6382 | 1.5925 | 1.75 | 1.76 | 1.77 | 1.78 | 1.89 |  |  | 16
| 1 | 4 | 1.85 | 1.756 | 1.729 | 1.9001 | 1.8353 | 1.79 | 1.9841 | 2.01 | 1.75 | 1.9 | 1.72 | 1.78 | 1.8 | 1.76 | 1.8193 | 1.86 | 1.7 | 1.8927 | 1.83 | 1.74 | 1.79 | 2.6623 | 1.6336 | 1.649 | 1.73 | 1.81 | 1.78 | 1.77 | 1.86 |  |  | 1.8113
| 1 | 5 | 1.86 | 1.748 | 1.7165 | 1.9602 | 1.8369 | 1.81 | 2.0082 | 2 | 1.81 | 1.92 | 1.71 | 1.79 | 1.79 | 1.77 | 1.7554 | 1.94 | 1.69 | 1.8924 | 1.91 | 1.75 | 1.81 | 2.1994 | 1.6477 | 1.572 | 1.74 | 1.77 | 1.8 | 1.8 | 1.88 |  |  | 103
| 1 | 6 | 1.87 | 1.741 | 1.7943 | 1.9507 | 1.8457 | 1.94 | 2.0022 | 1.96 | 1.84 | 1.92 | 1.72 | 1.77 | 1.8 | 1.78 | 1.7534 | 1.92 | 1.68 | 1.9018 | 1.79 | 1.78 | 1.8 | 2.5973 | 1.6507 | 1.582 | 1.75 | 1.78 | 1.78 | 1.78 | 1.85
| Mean |  | 1.8617 | 1.758 | 1.762 | 1.921 | 1.8378 | 1.885 | 2.0042 | 1.9817 | 1.8133 | 1.8867 | 1.7267 | 1.785 | 1.7983 | 1.7617 | 1.7973 | 1.8783 | 1.7067 | 1.8974 | 1.8583 | 1.7767 | 1.8067 | 2.4616 | 1.6132 | 1.6158 | 1.745 | 1.7867 | 1.7833 | 1.7833 | 1.8717
| Median |  | 1.865 | 1.752 | 1.7639 | 1.909 | 1.8361 | 1.9 | 2.0052 | 1.99 | 1.82 | 1.895 | 1.725 | 1.785 | 1.8 | 1.76 | 1.7873 | 1.86 | 1.695 | 1.8959 | 1.85 | 1.765 | 1.805 | 2.5513 | 1.6359 | 1.6182 | 1.745 | 1.775 | 1.78 | 1.78 | 1.875
| Std Dev. |  | 0.0098319 | 0.020659 | 0.034858 | 0.027404 | 0.0087546 | 0.071484 | 0.015564 | 0.026394 | 0.033267 | 0.03559 | 0.013663 | 0.010488 | 0.013292 | 0.01169 | 0.053095 | 0.042151 | 0.032042 | 0.02253 | 0.050761 | 0.040825 | 0.013663 | 0.27104 | 0.045875 | 0.037623 | 0.010488 | 0.027325 | 0.013663 | 0.013663 | 0.01472
| Rel.Std.Dev. |  | 0.5281% | 1.175% | 1.978% | 1.427% | 0.4764% | 3.792% | 0.7766% | 1.332% | 1.835% | 1.886% | 0.7913% | 0.5876% | 0.7391% | 0.6636% | 2.954% | 2.244% | 1.877% | 1.187% | 2.732% | 2.298% | 0.7562% | 11.01% | 2.844% | 2.328% | 0.601% | 1.529% | 0.7661% | 0.7661% | 0.7864%
| PDM3 |  | 2.781% | -2.942% | -2.719% | 6.056% | 1.465% | 4.07% | 10.65% | 9.407% | 0.1129% | 4.162% | -4.672% | -1.451% | -0.7152% | -2.74% | -0.7745% | 3.702% | -5.776% | 4.754% | 2.597% | -1.911% | -0.2551% | 35.9% | -10.94% | -10.79% | -3.66% | -1.359% | -1.543% | -1.543% | 3.333%
| Z-Score (Absolute) |  | 0.71 | 0.45 | 0.41 | 1.38 | 0.44 | 0.97 | 2.31 | 2.06 | 0.17 | 0.99 | 0.81 | 0.15 | 0 | 0.41 | 0.01 | 0.9 | 1.03 | 1.11 | 0.67 | 0.24 | 0.09 | 7.45 | 2.08 | 2.05 | 0.6 | 0.13 | 0.17 | 0.17 | 0.82
| Analytical results for Mn in OREAS 233b (Certified Value 0.081 wt.%) | [32 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.0843 | 0.08281 | 0.071957 | 0.085531 | 0.082347 | 0.0869 | 0.0914 | 0.0856 | 0.0838 | 0.0755 | 0.0947 | 0.0769 | 0.0778 | 0.0769 | 0.076537 | 0.0817 | 0.0746 | 0.084416 | 0.086 | 0.0767 | 0.0832 | 0.076073 | 0.080461 | 0.0705 | 0.0754 | 0.0766 | 0.0828 | 0.0745 | 0.0807 |  |  | 1
| 1 | 2 | 0.0854 | 0.0824 | 0.070594 | 0.085988 | 0.081608 | 0.089 | 0.0898 | 0.0839 | 0.0831 | 0.0762 | 0.0985 | 0.0772 | 0.0796 | 0.0779 | 0.079739 | 0.081 | 0.0783 | 0.086427 | 0.0823 | 0.0798 | 0.0827 | 0.074462 | 0.081149 | 0.0705 | 0.0744 | 0.078 | 0.0858 | 0.0754 | 0.0832 |  |  | 23
| 1 | 3 | 0.0857 | 0.08415 | 0.070207 | 0.086013 | 0.088342 | 0.0864 | 0.0914 | 0.0861 | 0.0844 | 0.0757 | 0.0918 | 0.0766 | 0.0779 | 0.0783 | 0.082335 | 0.0797 | 0.0771 | 0.087845 | 0.0854 | 0.0783 | 0.0819 | 0.076123 | 0.085618 | 0.068 | 0.0749 | 0.0752 | 0.0868 | 0.0747 | 0.0833 |  |  | 16
| 1 | 4 | 0.0857 | 0.08329 | 0.07008 | 0.085989 | 0.081599 | 0.0853 | 0.0891 | 0.088 | 0.0802 | 0.075 | 0.0964 | 0.0769 | 0.0786 | 0.0783 | 0.079566 | 0.0809 | 0.0752 | 0.085857 | 0.0808 | 0.0758 | 0.0818 | 0.076259 | 0.086314 | 0.07 | 0.0751 | 0.0774 | 0.0878 | 0.0741 | 0.0819 |  |  | 0.080881
| 1 | 5 | 0.0858 | 0.08319 | 0.070466 | 0.086262 | 0.080931 | 0.0867 | 0.0898 | 0.0867 | 0.0825 | 0.0752 | 0.0956 | 0.0765 | 0.078 | 0.0789 | 0.076296 | 0.085 | 0.0748 | 0.085218 | 0.0847 | 0.077 | 0.0819 | 0.07563 | 0.085443 | 0.067 | 0.0735 | 0.0755 | 0.0852 | 0.0751 | 0.0825 |  |  | 104
| 1 | 6 | 0.0859 | 0.08273 | 0.071002 | 0.086299 | 0.082646 | 0.0885 | 0.0906 | 0.0852 | 0.0832 | 0.0762 | 0.0951 | 0.0757 | 0.0785 | 0.0785 | 0.076466 | 0.0843 | 0.0744 | 0.085843 | 0.082 | 0.0771 | 0.0816 | 0.076734 | 0.085265 | 0.0675 | 0.0746 | 0.0756 | 0.0836 | 0.0744 | 0.0813
| Mean |  | 0.085467 | 0.083095 | 0.070718 | 0.086014 | 0.082912 | 0.087133 | 0.09035 | 0.085917 | 0.082867 | 0.075633 | 0.09535 | 0.076633 | 0.0784 | 0.078133 | 0.07849 | 0.0821 | 0.075733 | 0.085934 | 0.083533 | 0.07745 | 0.082183 | 0.07588 | 0.084042 | 0.068917 | 0.07465 | 0.076383 | 0.085333 | 0.0747 | 0.08215
| Median |  | 0.0857 | 0.083 | 0.07053 | 0.086001 | 0.081977 | 0.0868 | 0.0902 | 0.08585 | 0.08315 | 0.0756 | 0.09535 | 0.07675 | 0.07825 | 0.0783 | 0.078051 | 0.08135 | 0.075 | 0.08585 | 0.0835 | 0.07705 | 0.0819 | 0.076098 | 0.085354 | 0.069 | 0.07475 | 0.0761 | 0.0855 | 0.0746 | 0.0822
| Std Dev. |  | 0.00059554 | 0.00060952 | 0.00068743 | 0.00027471 | 0.0027287 | 0.0013779 | 0.00094181 | 0.0013906 | 0.0014583 | 0.00050067 | 0.0021989 | 0.00052026 | 0.00067231 | 0.00068605 | 0.0024588 | 0.002089 | 0.0015921 | 0.0011599 | 0.0021106 | 0.0014039 | 0.00062423 | 0.00078006 | 0.0025415 | 0.0015943 | 0.00066558 | 0.0011392 | 0.0018917 | 0.00047749 | 0.0010426
| Rel.Std.Dev. |  | 0.6968% | 0.7335% | 0.9721% | 0.3194% | 3.291% | 1.581% | 1.042% | 1.619% | 1.76% | 0.662% | 2.306% | 0.6789% | 0.8575% | 0.8781% | 3.133% | 2.544% | 2.102% | 1.35% | 2.527% | 1.813% | 0.7596% | 1.028% | 3.024% | 2.313% | 0.8916% | 1.491% | 2.217% | 0.6392% | 1.269%
| PDM3 |  | 5.67% | 2.737% | -12.57% | 6.346% | 2.511% | 7.73% | 11.71% | 6.226% | 2.455% | -6.488% | 17.89% | -5.252% | -3.068% | -3.397% | -2.957% | 1.507% | -6.365% | 6.248% | 3.279% | -4.242% | 1.61% | -6.183% | 3.908% | -14.79% | -7.704% | -5.561% | 5.505% | -7.642% | 1.569%
| Z-Score (Absolute) |  | 0.58 | 0.16 | 2 | 0.67 | 0.13 | 0.87 | 1.43 | 0.66 | 0.13 | 1.14 | 2.3 | 0.96 | 0.65 | 0.7 | 0.64 | 0.01 | 1.12 | 0.66 | 0.24 | 0.82 | 0.01 | 1.09 | 0.33 | 2.31 | 1.31 | 1.01 | 0.56 | 1.3 | 0
| Analytical results for Mo in OREAS 233b (Certified Value 2.26 ppm) | [32 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 3
| 1 | 1 | 2.22 | 2.23 | 2.58 | 1.848 | 2.4175 | 2.28 | 2.31 | 2.3 | 2.8 | 2 | 2.3 | 2.24 | 2.2 | 2.18 | < 10 | 2.07 | < 2 | < 5 | 2.1 | 1.97 | 2.33 | 2.16 | 2 | 2.2889 | 2.25 | 2.262 | 2.31 | 2.33 | 2.27 | 2.26 |  | 1
| 1 | 2 | 2.27 | 2.25 | 2.41 | 1.821 | 2.338 | 2.23 | 2.04 | 2.2 | 2.4 | 2 | 2.7 | 2.29 | 2.24 | 2.2 | < 10 | 2.03 | < 2 | < 5 | 2 | 2.07 | 2.32 | 2.62 | 2 | 2.2661 | 2.22 | 2.229 | 2.36 | 2.44 | 2.3 | 2.28 |  | 24
| 1 | 3 | 2.28 | 2.27 | 2.17 | 1.8378 | 2.556 | 2.1 | 2.02 | 2.3 | 2.6 | 2 | 2.2 | 2.28 | 2.22 | 2.27 | < 10 | 2.02 | < 2 | < 5 | 2 | 2.01 | 2.26 | 1.66 | 2 | 2.3141 | 2.24 | 2.103 | 2.3 | 2.45 | 2.31 | 2.23 |  | 16
| 1 | 4 | 2.31 | 2.3 | 2.31 | 1.7791 | 2.4276 | 2.14 | 2.08 | 2.5 | 2.4 | 2 | 2.4 | 2.12 | 2.2 | 2.2 | < 10 | 2.08 | < 2 | < 5 | 1.9 | 2 | 2.28 | 2.45 | 2 | 2.2635 | 2.25 | 2.186 | 2.4 | 2.48 | 2.25 | 2.25 |  | 2.2599
| 1 | 5 | 2.25 | 2.25 | 2.08 | 1.7803 | 2.1988 | 2.15 | 2.03 | 2.3 | 2.2 | 2 | 2.2 | 2.2 | 2.23 | 2.17 | < 10 | 2.21 | < 2 | < 5 | 1.9 | 1.99 | 2.46 | 1.93 | 2 | 2.1607 | 2.24 | 2.228 | 2.32 | 2.41 | 2.25 | 2.31 |  | 105
| 1 | 6 | 2.25 | 2.25 | 2.36 | 1.8483 | 2.3339 | 2.21 | 2.03 | 2.2 | 2.2 | 2 | 2.3 | 2.2 | 2.21 | 2.32 | < 10 | 2.15 | < 2 | < 5 | 2.1 | 2.01 | 2.27 | 1.64 | 2 | 2.2902 | 2.24 | 2.053 | 2.47 | 2.34 | 2.24 | 2.21
| Mean |  | 2.2633 | 2.2583 | 2.3183 | 1.8191 | 2.3786 | 2.185 | 2.085 | 2.3 | 2.4333 | 2 | 2.35 | 2.2217 | 2.2167 | 2.2233 |  | 2.0933 |  |  | 2 | 2.0083 | 2.32 | 2.0767 | 2 | 2.2639 | 2.24 | 2.1768 | 2.36 | 2.4083 | 2.27 | 2.2567
| Median |  | 2.26 | 2.25 | 2.335 | 1.8294 | 2.3777 | 2.18 | 2.035 | 2.3 | 2.4 | 2 | 2.3 | 2.22 | 2.215 | 2.2 |  | 2.075 |  |  | 2 | 2.005 | 2.3 | 2.045 | 2 | 2.2775 | 2.24 | 2.207 | 2.34 | 2.425 | 2.26 | 2.255
| Std Dev. |  | 0.030768 | 0.024014 | 0.17747 | 0.032096 | 0.11944 | 0.066558 | 0.11221 | 0.10954 | 0.23381 | 0 | 0.18708 | 0.062743 | 0.01633 | 0.058878 |  | 0.073394 |  |  | 0.089443 | 0.033714 | 0.074027 | 0.40678 | 0 | 0.053824 | 0.010954 | 0.081808 | 0.065422 | 0.061128 | 0.028983 | 0.03559
| Rel.Std.Dev. |  | 1.359% | 1.063% | 7.655% | 1.764% | 5.021% | 3.046% | 5.382% | 4.763% | 9.609% | 0% | 7.961% | 2.824% | 0.7367% | 2.648% |  | 3.506% |  |  | 4.472% | 1.679% | 3.191% | 19.59% | 0% | 2.377% | 0.489% | 3.758% | 2.772% | 2.538% | 1.277% | 1.577%
| PDM3 |  | 0.1535% | -0.06775% | 2.587% | -19.5% | 5.255% | -3.313% | -7.738% | 1.776% | 7.676% | -11.5% | 3.989% | -1.69% | -1.912% | -1.617% |  | -7.369% |  |  | -11.5% | -11.13% | 2.661% | -8.107% | -11.5% | 0.1794% | -0.879% | -3.674% | 4.431% | 6.57% | 0.4485% | -0.1415%
| Z-Score (Absolute) |  | 0.05 | 0.01 | 0.47 | 3.41 | 0.94 | 0.56 | 1.34 | 0.33 | 1.37 | NA | 0.72 | 0.28 | 0.32 | 0.27 | 21.36 | 1.28 | 9.79 | 1.89 | 2.01 | 1.94 | 0.49 | 1.41 | NA | 0.05 | 0.14 | 0.63 | 0.8 | 1.17 | 0.1 | 0.01
Results from laboratories 1.12 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 233b (Certified Value 0.221 wt.%) | [32 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.23 | 0.221 | 1.1099 | 0.20149 | 0.21669 | 0.208 | 0.237 | 0.2 | 0.21 | 0.21 | 0.22 | 0.22 | 0.21 | 0.23 | 0.27466 | 0.22 | 0.21 | 0.21575 | 0.22 | 0.238 | 0.22 | 0.25 | 0.234 | 0.2095 | 0.21 | 0.2 | 0.27 | 0.23 | 0.23 |  |  | 1
| 1 | 2 | 0.23 | 0.22 | 1.113 | 0.21333 | 0.21411 | 0.211 | 0.23 | 0.2 | 0.21 | 0.21 | 0.22 | 0.21 | 0.21 | 0.24 | 0.28242 | 0.22 | 0.22 | 0.21787 | 0.22 | 0.257 | 0.22 | 0.24 | 0.23406 | 0.2125 | 0.21 | 0.21 | 0.27 | 0.23 | 0.23 |  |  | 0
| 1 | 3 | 0.24 | 0.225 | 1.1701 | 0.20792 | 0.23204 | 0.198 | 0.237 | 0.19 | 0.21 | 0.2 | 0.22 | 0.21 | 0.21 | 0.24 | 0.29015 | 0.22 | 0.21 | 0.22388 | 0.22 | 0.251 | 0.22 | 0.25 | 0.2483 | 0.203 | 0.21 | 0.2 | 0.26 | 0.23 | 0.24 |  |  | 16
| 1 | 4 | 0.23 | 0.222 | 1.0697 | 0.20941 | 0.21595 | 0.195 | 0.223 | 0.22 | 0.2 | 0.22 | 0.22 | 0.21 | 0.21 | 0.24 | 0.27974 | 0.22 | 0.21 | 0.21848 | 0.22 | 0.243 | 0.22 | 0.26 | 0.24913 | 0.2105 | 0.21 | 0.21 | 0.27 | 0.23 | 0.24 |  |  | 0.22117
| 1 | 5 | 0.24 | 0.219 | 1.047 | 0.20281 | 0.21504 | 0.208 | 0.23 | 0.2 | 0.21 | 0.22 | 0.21 | 0.21 | 0.21 | 0.24 | 0.27278 | 0.23 | 0.2 | 0.21899 | 0.22 | 0.247 | 0.23 | 0.26 | 0.24715 | 0.207 | 0.21 | 0.21 | 0.27 | 0.23 | 0.25 |  |  | 106
| 1 | 6 | 0.24 | 0.22 | 1.1979 | 0.20612 | 0.22017 | 0.209 | 0.237 | 0.2 | 0.21 | 0.22 | 0.22 | 0.21 | 0.21 | 0.24 | 0.27538 | 0.23 | 0.21 | 0.21901 | 0.22 | 0.253 | 0.22 | 0.26 | 0.25159 | 0.2035 | 0.21 | 0.2 | 0.27 | 0.23 | 0.25
| Mean |  | 0.235 | 0.22117 | 1.1179 | 0.20685 | 0.219 | 0.20483 | 0.23233 | 0.20167 | 0.20833 | 0.21333 | 0.21833 | 0.21167 | 0.21 | 0.23833 | 0.27919 | 0.22333 | 0.21 | 0.219 | 0.22 | 0.24817 | 0.22167 | 0.25333 | 0.24404 | 0.20767 | 0.21 | 0.205 | 0.26833 | 0.23 | 0.24
| Median |  | 0.235 | 0.2205 | 1.1115 | 0.20702 | 0.21632 | 0.208 | 0.2335 | 0.2 | 0.21 | 0.215 | 0.22 | 0.21 | 0.21 | 0.24 | 0.27756 | 0.22 | 0.21 | 0.21874 | 0.22 | 0.249 | 0.22 | 0.255 | 0.24773 | 0.20825 | 0.21 | 0.205 | 0.27 | 0.23 | 0.24
| Std Dev. |  | 0.0054772 | 0.002137 | 0.057542 | 0.0043652 | 0.0067183 | 0.0066156 | 0.0057155 | 0.0098319 | 0.0040825 | 0.008165 | 0.0040825 | 0.0040825 | 0 | 0.0040825 | 0.0064337 | 0.005164 | 0.0063246 | 0.0026795 | 0 | 0.0069402 | 0.0040825 | 0.008165 | 0.0078884 | 0.0038557 | 0 | 0.0054772 | 0.0040825 | 0 | 0.0089443
| Rel.Std.Dev. |  | 2.331% | 0.9662% | 5.147% | 2.11% | 3.068% | 3.23% | 2.46% | 4.875% | 1.96% | 3.827% | 1.87% | 1.929% | 0% | 1.713% | 2.304% | 2.312% | 3.012% | 1.224% | 0% | 2.797% | 1.842% | 3.223% | 3.232% | 1.857% | 0% | 2.672% | 1.521% | 0% | 3.727%
| PDM3 |  | 6.251% | -0.003751% | 405.5% | -6.478% | -0.9837% | -7.389% | 5.045% | -8.82% | -5.806% | -3.545% | -1.285% | -4.299% | -5.053% | 7.758% | 26.23% | 0.9759% | -5.053% | -0.9853% | -0.5312% | 12.2% | 0.2223% | 14.54% | 10.34% | -6.108% | -5.053% | -7.313% | 21.32% | 3.99% | 8.511%
| Z-Score (Absolute) |  | 0.82 | 0.06 | 49.09 | 0.72 | 0.05 | 0.83 | 0.68 | 1 | 0.64 | 0.36 | 0.09 | 0.46 | 0.55 | 1 | 3.24 | 0.18 | 0.55 | 0.05 | 0 | 1.54 | 0.09 | 1.82 | 1.31 | 0.67 | 0.55 | 0.82 | 2.64 | 0.55 | 1.09
| Analytical results for Nb in OREAS 233b (Indicative Value 0.11 ppm) | [32 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | < 0.05 | 0.13 | 34.02 | 0.14226 | 0.1 | 0.38 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.09 | 0.14 | 0.09 | 0.06 | < 10 | < 0.2 | < 1 | 0.11 | 0.59 | 0.19 | 0.07 | 0.13 | [9 empty] | 1
| 1 | 2 | < 0.05 | 0.14 | 35.36 | 0.15204 | < 0.1 | 0.38 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.09 | 0.15 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.12 | 0.61 | 0.21 | 0.08 | 0.14 | [9 empty] | 7
| 1 | 3 | < 0.05 | 0.14 | 33.53 | 0.2218 | < 0.1 | 0.36 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.09 | 0.14 | 0.09 | 0.06 | < 10 | < 0.2 | < 1 | 0.1 | 0.47 | 0.21 | 0.07 | 0.13 | [9 empty] | 16
| 1 | 4 | < 0.05 | 0.13 | 36.09 | 0.15785 | < 0.1 | 0.37 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.09 | 0.14 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.36 | 0.2 | 0.08 | 0.12 | [9 empty] | 0.11493
| 1 | 5 | < 0.05 | 0.13 | 34.38 | 0.13874 | < 0.1 | 0.35 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.1 | 0.14 | 0.1 | 0.06 | < 10 | < 0.2 | < 1 | 0.11 | 0.35 | 0.2 | 0.07 | 0.13 | [9 empty] | 13
| 1 | 6 | < 0.05 | 0.13 | 33.44 | 0.14676 | < 0.1 | 0.36 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.09 | 0.14 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.1 | 0.28 | 0.19 | 0.07 | 0.14
| Mean |  | 1.02: 0.13333; 1.03: 34.47; 1.04: 0.15991; 1.05: 0.1; 1.07: 0.36667; 1.10: 0.095; 1.13: 0.091667; 1.14: 0.14167; 1.15: 0.091667; 1.16: 0.065; 1.21: 0.10833; 1.22: 0.44333; 1.23: 0.2; 1.24: 0.073333; 1.25: 0.13167
| Median |  | 1.02: 0.13; 1.03: 34.2; 1.04: 0.1494; 1.05: 0.1; 1.07: 0.365; 1.10: 0.095; 1.13: 0.09; 1.14: 0.14; 1.15: 0.09; 1.16: 0.065; 1.21: 0.11; 1.22: 0.415; 1.23: 0.2; 1.24: 0.07; 1.25: 0.13
| Std Dev. |  | 1.02: 0.005164; 1.03: 1.0556; 1.04: 0.031079; 1.07: 0.012111; 1.10: 0.0054772; 1.13: 0.0040825; 1.14: 0.0040825; 1.15: 0.0040825; 1.16: 0.0054772; 1.21: 0.0075277; 1.22: 0.13589; 1.23: 0.0089443; 1.24: 0.005164; 1.25: 0.0075277
| Rel.Std.Dev. |  | 1.02: 3.873%; 1.03: 3.062%; 1.04: 19.44%; 1.07: 3.303%; 1.10: 5.766%; 1.13: 4.454%; 1.14: 2.882%; 1.15: 4.454%; 1.16: 8.427%; 1.21: 6.949%; 1.22: 30.65%; 1.23: 4.472%; 1.24: 7.042%; 1.25: 5.717%
| PDM3 |  | 1.02: 16.01%; 1.03: 29891%; 1.04: 39.13%; 1.05: -12.99%; 1.07: 219%; 1.10: -17.34%; 1.13: -20.24%; 1.14: 23.26%; 1.15: -20.24%; 1.16: -43.45%; 1.21: -5.742%; 1.22: 285.7%; 1.23: 74.01%; 1.24: -36.19%; 1.25: 14.56%
| Z-Score (Absolute) |  | 1.26 | 0.18 | 455.66 | 0.53 | 0.82 | 3.27 | 0.27 | 0.33 | 31.57 | 0.93 | 0.38 | 0.29 | 0.38 | 0.73 | 64.73 | 0.27 | 5.04 | 0.15 | 4.29 | 1.06 | 0.62 | 0.15
| Analytical results for Nd in OREAS 233b (Certified Value 5.63 ppm) | [32 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.27 |  |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.07: AR*MS |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g |  |  | 3
| 1 | 1 | 1.01: 6.098; 1.02: 5.99; 1.03: 5.1097; 1.04: 6.31; 1.05: 5.96; 1.07: 5.5 |  |  | 1
| 1 | 2 | 1.01: 5.999; 1.02: 5.584; 1.03: 5.1475; 1.04: 6.22; 1.05: 6.07; 1.07: 5.5 |  |  | 26
| 1 | 3 | 1.01: 6.145; 1.02: 5.231; 1.03: 5.1171; 1.04: 5.86; 1.05: 5.9; 1.07: 5 |  |  | 16
| 1 | 4 | 1.01: 6.128; 1.02: 5.228; 1.03: 5.0126; 1.04: 5.61; 1.05: 6.01; 1.07: 5 |  |  | 5.6347
| 1 | 5 | 1.01: 6.093; 1.02: 5.403; 1.03: 5.0388; 1.04: 5.88; 1.05: 5.87; 1.07: 5 |  |  | 107
| 1 | 6 | 1.01: 6.058; 1.02: 5.682; 1.03: 5.233; 1.04: 5.84; 1.05: 6.02; 1.07: 5
| Mean |  | 1.01: 6.0868; 1.02: 5.5197; 1.03: 5.1098; 1.04: 5.9533; 1.05: 5.9717; 1.07: 5.1667
| Median |  | 1.01: 6.0955; 1.02: 5.4935; 1.03: 5.1134; 1.04: 5.87; 1.05: 5.985; 1.07: 5
| Std Dev. |  | 1.01: 0.052511; 1.02: 0.29452; 1.03: 0.078918; 1.04: 0.26197; 1.05: 0.076267; 1.07: 0.2582
| Rel.Std.Dev. |  | 1.01: 0.8627%; 1.02: 5.336%; 1.03: 1.544%; 1.04: 4.4%; 1.05: 1.277%; 1.07: 4.997%
| PDM3 |  | 1.01: 8.025%; 1.02: -2.041%; 1.03: -9.315%; 1.04: 5.656%; 1.05: 5.981%; 1.07: -8.306%
| Z-Score (Absolute) |  | 1.01: 0.81; 1.02: 0.5; 1.03: 1.44; 1.04: 0.5; 1.05: 0.54; 1.07: 1.31
| Analytical results for Ni in OREAS 233b (Certified Value 68 ppm) | [32 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 1
| 1 | 1 | 67.3 | 68.3 | 58.88 | 64.91 | 70.6632 | 69.4 | 68.3 | 75 | 71.1 | 61 | 63.2 | 69.6 | 67.8 | 67.4 | 66.019 | 68.4 | 59 | 71.1724 | 70 | 68 | 68 | 63.7 | 63.7192 | 64 | 71.8961 | 67.2 | 65.3 | 70.3 | 63.7 | 69.1 |  | 1
| 1 | 2 | 68.4 | 67.7 | 57.8 | 66.33 | 70.4078 | 68.6 | 68.9 | 74 | 70.5 | 60 | 69.3 | 71 | 69.9 | 70.2 | 67.476 | 67.5 | 61 | 71.6974 | 65 | 74.2 | 67 | 64.69 | 63.5996 | 63 | 70.4012 | 67.9 | 66.5 | 74.2 | 64.1 | 69.4 |  | 27
| 1 | 3 | 68.3 | 68.7 | 58.34 | 66.35 | 76.0889 | 66.7 | 68.6 | 74 | 69.3 | 60 | 61.7 | 71.4 | 68.6 | 69.3 | 68.827 | 67.4 | 60 | 73.2365 | 71 | 70.4 | 67 | 65.07 | 68.034 | 62 | 70.3449 | 67.6 | 65.8 | 73.9 | 63.9 | 68.5 |  | 16
| 1 | 4 | 67.9 | 68.7 | 58.8 | 65.23 | 70.387 | 65.4 | 71.3 | 76 | 69.3 | 61 | 62.5 | 69.3 | 67.6 | 69.6 | 66.982 | 67.6 | 59 | 72.1804 | 66 | 68.6 | 67 | 65.63 | 67.5832 | 62 | 70.4296 | 67.9 | 67.7 | 74.1 | 63.1 | 68.3 |  | 67.6337
| 1 | 5 | 68.7 | 67.5 | 58.86 | 67.42 | 69.7882 | 66.5 | 69.4 | 76 | 69 | 62 | 62.6 | 71.1 | 68.3 | 69.8 | 64.909 | 71.8 | 59 | 71.9047 | 68 | 68.9 | 67 | 64.8 | 67.1692 | 60 | 71.0592 | 67.9 | 65.5 | 72.2 | 62.7 | 68.6 |  | 108
| 1 | 6 | 68.9 | 66.8 | 57.58 | 68.03 | 71.3369 | 68.9 | 72.6 | 76 | 69 | 62 | 62.2 | 69.5 | 68.1 | 71.1 | 66.72 | 70.3 | 58 | 73.4168 | 66 | 69.5 | 66 | 66.68 | 65.7432 | 59 | 71.4866 | 67.3 | 67.3 | 72.8 | 63 | 67.2
| Mean |  | 68.25 | 67.95 | 58.3767 | 66.3783 | 71.4453 | 67.5833 | 69.85 | 75.1667 | 69.7 | 61 | 63.5833 | 70.3167 | 68.3833 | 69.5667 | 66.8222 | 68.8333 | 59.3333 | 72.2681 | 67.6667 | 69.9333 | 67 | 65.095 | 65.9747 | 61.6667 | 70.9363 | 67.6333 | 66.35 | 72.9167 | 63.4167 | 68.5167
| Median |  | 68.35 | 68 | 58.57 | 66.34 | 70.5355 | 67.65 | 69.15 | 75.5 | 69.3 | 61 | 62.55 | 70.3 | 68.2 | 69.7 | 66.851 | 68 | 59 | 72.0426 | 67 | 69.2 | 67 | 64.935 | 66.4562 | 62 | 70.7444 | 67.75 | 66.15 | 73.35 | 63.4 | 68.55
| Std Dev. |  | 0.57879 | 0.75299 | 0.57179 | 1.2075 | 2.3293 | 1.5993 | 1.7167 | 0.98319 | 0.88318 | 0.89443 | 2.8435 | 0.94534 | 0.82321 | 1.2307 | 1.3258 | 1.8162 | 1.0328 | 0.88583 | 2.4221 | 2.2447 | 0.63246 | 0.99961 | 1.9512 | 1.8619 | 0.65298 | 0.32042 | 0.98742 | 1.5092 | 0.56006 | 0.76267
| Rel.Std.Dev. |  | 0.848% | 1.108% | 0.9795% | 1.819% | 3.26% | 2.366% | 2.458% | 1.308% | 1.267% | 1.466% | 4.472% | 1.344% | 1.204% | 1.769% | 1.984% | 2.639% | 1.741% | 1.226% | 3.579% | 3.21% | 0.944% | 1.536% | 2.958% | 3.019% | 0.9205% | 0.4738% | 1.488% | 2.07% | 0.8831% | 1.113%
| PDM3 |  | 0.9112% | 0.4676% | -13.69% | -1.856% | 5.636% | -0.07454% | 3.277% | 11.14% | 3.055% | -9.808% | -5.989% | 3.967% | 1.108% | 2.858% | -1.2% | 1.774% | -12.27% | 6.852% | 0.04867% | 3.4% | -0.937% | -3.754% | -2.453% | -8.823% | 4.883% | -0.0006122% | -1.898% | 7.811% | -6.235% | 1.305%
| Z-Score (Absolute) |  | 0.15 | 0.05 | 3.23 | 0.49 | 1.25 | 0.08 | 0.7 | 2.52 | 0.65 | 2.33 | 1.45 | 0.86 | 0.2 | 0.6 | 0.34 | 0.35 | 2.91 | 1.53 | 0.05 | 0.73 | 0.28 | 0.93 | 0.63 | 2.11 | 1.07 | 0.06 | 0.5 | 1.75 | 1.51 | 0.24
| Analytical results for P in OREAS 233b (Certified Value 0.042 wt.%) | [32 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.043 | 0.0425 | 0.035531 | 0.041889 | 0.039 | 0.041 | 0.0434 | 0.0399 | 0.0422 | 0.044 | 0.043 | 0.042 | 0.041 | 0.044557 | 0.043 | 0.038 | 0.044845 | 0.041 | 0.042 | 0.04 | 0.0425 | 0.041 | 0.041 | 0.0446 | 0.0401 | 0.041 | [5 empty] | 1
| 1 | 2 | 0.0437 | 0.0415 | 0.034621 | 0.041437 | 0.04 | 0.0406 | 0.0424 | 0.0415 | 0.0425 | 0.054 | 0.043 | 0.043 | 0.04 | 0.04389 | 0.042 | 0.0392 | 0.045262 | 0.043 | 0.043 | 0.04 | 0.0425 | 0.04 | 0.042 | 0.0437 | 0.0403 | 0.042 | [5 empty] | 0
| 1 | 3 | 0.0432 | 0.0433 | 0.034394 | 0.045354 | 0.039 | 0.0401 | 0.0432 | 0.0401 | 0.0413 | 0.043 | 0.042 | 0.043 | 0.04 | 0.048138 | 0.042 | 0.0388 | 0.04629 | 0.043 | 0.041 | 0.04 | 0.041 | 0.041 | 0.042 | 0.0433 | 0.0401 | 0.042 | [5 empty] | 16
| 1 | 4 | 0.0433 | 0.0431 | 0.033948 | 0.041965 | 0.038 | 0.0393 | 0.0439 | 0.0381 | 0.0428 | 0.044 | 0.042 | 0.043 | 0.04 | 0.046098 | 0.042 | 0.0379 | 0.045584 | 0.042 | 0.041 | 0.04 | 0.042 | 0.04 | 0.042 | 0.0435 | 0.0398 | 0.042 | [5 empty] | 0.0419
| 1 | 5 | 0.0438 | 0.0424 | 0.034761 | 0.042198 | 0.039 | 0.041 | 0.0436 | 0.0436 | 0.0429 | 0.046 | 0.042 | 0.043 | 0.041 | 0.043135 | 0.044 | 0.0381 | 0.045473 | 0.044 | 0.041 | 0.04 | 0.0415 | 0.04 | 0.041 | 0.0427 | 0.0405 | 0.042 | [5 empty] | 109
| 1 | 6 | 0.0437 | 0.042 | 0.034503 | 0.042265 | 0.039 | 0.0401 | 0.0426 | 0.0404 | 0.0425 | 0.046 | 0.042 | 0.043 | 0.04 | 0.044237 | 0.044 | 0.038 | 0.046064 | 0.041 | 0.041 | 0.04 | 0.0405 | 0.041 | 0.042 | 0.0437 | 0.04 | 0.041
| Mean |  | 0.04345 | 0.042467 | 0.034626 | 0.042518 | 0.039 | 0.04035 | 0.043183 | 0.0406 | 0.042367 | 0.046167 | 0.042333 | 0.042833 | 0.040333 | 0.045009 | 0.042833 | 0.038333 | 0.045586 | 0.042333 | 0.0415 | 0.04 | 0.041667 | 0.0405 | 0.041667 | 0.043583 | 0.040133 | 0.041667
| Median |  | 0.0435 | 0.04245 | 0.034562 | 0.042082 | 0.039 | 0.04035 | 0.0433 | 0.04025 | 0.0425 | 0.045 | 0.042 | 0.043 | 0.04 | 0.044397 | 0.0425 | 0.03805 | 0.045528 | 0.0425 | 0.041 | 0.04 | 0.04175 | 0.0405 | 0.042 | 0.0436 | 0.0401 | 0.042
| Std Dev. |  | 0.00032711 | 0.00067132 | 0.00052267 | 0.0014197 | 0.00063246 | 0.00065345 | 0.00058109 | 0.0018352 | 0.0005785 | 0.0040208 | 0.0005164 | 0.00040825 | 0.0005164 | 0.0018189 | 0.00098319 | 0.00053541 | 0.00052737 | 0.0012111 | 0.00083666 | 0 | 0.0008165 | 0.00054772 | 0.0005164 | 0.00062102 | 0.00024221 | 0.0005164
| Rel.Std.Dev. |  | 0.7528% | 1.581% | 1.509% | 3.339% | 1.622% | 1.619% | 1.346% | 4.52% | 1.365% | 8.709% | 1.22% | 0.9531% | 1.28% | 4.041% | 2.295% | 1.397% | 1.157% | 2.861% | 2.016% | 0% | 1.96% | 1.352% | 1.239% | 1.425% | 0.6035% | 1.239%
| PDM3 |  | 3.7% | 1.353% | -17.36% | 1.476% | -6.921% | -3.699% | 3.063% | -3.102% | 1.114% | 10.18% | 1.035% | 2.228% | -3.739% | 7.421% | 2.228% | -8.512% | 8.798% | 1.035% | -0.9542% | -4.534% | -0.5565% | -3.341% | -0.5565% | 4.018% | -4.216% | -0.5565%
| Z-Score (Absolute) |  | 0.69 | 0.22 | 3.49 | 0.25 | 1.42 | 0.77 | 0.56 | 0.66 | 0.17 | 1.97 | 0.16 | 0.39 | 0.79 | 1.42 | 0.39 | 1.74 | 1.7 | 0.16 | 0.24 | 0.95 | 0.16 | 0.71 | 0.16 | 0.75 | 0.88 | 0.16
| Analytical results for Pb in OREAS 233b (Certified Value 23.6 ppm) | [32 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 2
| 1 | 1 | 24.1 | 23.1 | 22.4 | 22.1 | 23.8385 | 18.4 | 23.9 | 22.7 | 24.9 | 22 | 25.2 | 22.9 | 24.3 | 23.1 | 23.5 | 19 | 23.7232 | 22.5 | 21.9 | 23.9 | 26.28 | 21.2407 | 26 | 24.002 | 23.2 | 28.691 | 23.1 | 24 | 23.4 | 24.7 |  | 1
| 1 | 2 | 24.4 | 23.3 | 22.89 | 22.25 | 23.5964 | 17.2 | 24 | 22 | 24.9 | 23 | 27.9 | 22.9 | 24.9 | 23.3 | 23.5 | 20 | 22.5158 | 22.5 | 21.98 | 23.5 | 25.47 | 22.0827 | 26 | 24.146 | 23 | 29.187 | 23.9 | 24.2 | 23.9 | 24.8 |  | 28
| 1 | 3 | 24.3 | 23.3 | 23.94 | 22.03 | 25.9008 | 16.7 | 23.8 | 20.5 | 24.4 | 22 | 25 | 23.1 | 24.4 | 23.9 | 23.1 | 20 | 22.442 | 22.2 | 21.72 | 24.5 | 25.78 | 21.5148 | 25 | 24.0887 | 22.9 | 29.355 | 23.1 | 24 | 23 | 24.6 |  | 16
| 1 | 4 | 24 | 23.1 | 23.37 | 23.25 | 23.5959 | 16.7 | 24 | 22.8 | 23.6 | 23 | 25 | 23.2 | 24.2 | 23.8 | 23.4 | 20 | 21.8991 | 21 | 22.51 | 23.9 | 24.66 | 21.8868 | 26 | 24.3138 | 22.6 | 28.945 | 24 | 24.1 | 23.2 | 24.7 |  | 23.5553
| 1 | 5 | 24.4 | 23.1 | 23.48 | 23.22 | 22.8234 | 17.1 | 23.9 | 22.3 | 24.6 | 23 | 25.3 | 23.4 | 23.5 | 23.5 | 24.7 | 19 | 20.9442 | 23.6 | 21.43 | 24.6 | 26.37 | 22.7522 | 25 | 24.1961 | 23 | 29.296 | 24 | 23.8 | 23.2 | 24.5 |  | 110
| 1 | 6 | 24.7 | 23.2 | 23.63 | 23.35 | 23.8915 | 17.7 | 25.4 | 21.9 | 24.5 | 23 | 25.4 | 23.4 | 23.9 | 24.7 | 23.9 | 19 | 23.7403 | 20.6 | 21.56 | 24.4 | 25.42 | 21.5383 | 25 | 24.0459 | 22.9 | 31.626 | 23.6 | 23 | 23 | 24
| Mean |  | 24.3167 | 23.1833 | 23.285 | 22.7 | 23.9411 | 17.3 | 24.1667 | 22.0333 | 24.4833 | 22.6667 | 25.6333 | 23.15 | 24.2 | 23.7167 | 23.6833 | 19.5 | 22.5441 | 22.0667 | 21.85 | 24.1333 | 25.6633 | 21.8359 | 25.5 | 24.1321 | 22.9333 | 29.5167 | 23.6167 | 23.85 | 23.2833 | 24.55
| Median |  | 24.35 | 23.15 | 23.425 | 22.735 | 23.7174 | 17.15 | 23.95 | 22.15 | 24.55 | 23 | 25.25 | 23.15 | 24.25 | 23.65 | 23.5 | 19.5 | 22.4789 | 22.35 | 21.81 | 24.15 | 25.625 | 21.7126 | 25.5 | 24.1173 | 22.95 | 29.2415 | 23.75 | 24 | 23.2 | 24.65
| Std Dev. |  | 0.24833 | 0.098319 | 0.55363 | 0.63353 | 1.0335 | 0.65422 | 0.60882 | 0.83347 | 0.47924 | 0.5164 | 1.1219 | 0.22583 | 0.47329 | 0.56716 | 0.56006 | 0.54772 | 1.0778 | 1.0985 | 0.38273 | 0.43205 | 0.63191 | 0.53834 | 0.54772 | 0.11275 | 0.19664 | 1.0622 | 0.42622 | 0.43704 | 0.33714 | 0.2881
| Rel.Std.Dev. |  | 1.021% | 0.4241% | 2.378% | 2.791% | 4.317% | 3.782% | 2.519% | 3.783% | 1.957% | 2.278% | 4.377% | 0.9755% | 1.956% | 2.391% | 2.365% | 2.809% | 4.781% | 4.978% | 1.752% | 1.79% | 2.462% | 2.465% | 2.148% | 0.4672% | 0.8574% | 3.599% | 1.805% | 1.832% | 1.448% | 1.174%
| PDM3 |  | 3.232% | -1.579% | -1.147% | -3.631% | 1.638% | -26.56% | 2.595% | -6.461% | 3.94% | -3.773% | 8.822% | -1.721% | 2.737% | 0.6851% | 0.5436% | -17.22% | -4.293% | -6.32% | -7.24% | 2.454% | 8.949% | -7.299% | 8.256% | 2.449% | -2.64% | 25.31% | 0.2606% | 1.251% | -1.155% | 4.223%
| Z-Score (Absolute) |  | 0.58 | 0.41 | 0.32 | 0.83 | 0.25 | 5.52 | 0.45 | 1.41 | 0.73 | 0.86 | 1.73 | 0.44 | 0.48 | 0.06 | 0.03 | 3.61 | 0.96 | 1.38 | 1.57 | 0.42 | 1.75 | 1.58 | 1.61 | 0.42 | 0.62 | 5.1 | 0.03 | 0.17 | 0.32 | 0.78
| Analytical results for Pd in OREAS 233b (Indicative Value 17.6 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.11; 1.03: 1.22; 1.04: 1.27; 1.05: 1.28; 1.07: 1.34 |  | ppb
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS; 1.04: AR*MS; 1.05: AR*MS; 1.07: AR*MS |  | 1
|  |  | 1.01: 0.5g; 1.02: 01g; 1.03: 01g; 1.04: 25g; 1.05: 0.5g; 1.07: 0.5g |  | 1
| 1 | 1 | 1.01: 22; 1.02: 16; 1.03: 13; 1.04: < 50; 1.05: 22; 1.07: 42 |  | 1
| 1 | 2 | 1.01: 20; 1.02: 18; 1.03: 18; 1.04: < 50; 1.05: 20.55; 1.07: 30 |  | 1
| 1 | 3 | 1.01: 17; 1.02: 17; 1.03: 15; 1.04: < 50; 1.05: 28.28; 1.07: 30 |  | 16
| 1 | 4 | 1.01: 24; 1.02: 21; 1.03: 13; 1.04: < 50; 1.05: < 20; 1.07: 34 |  | 17.6179
| 1 | 5 | 1.01: 21; 1.02: 15; 1.03: 17; 1.04: < 50; 1.05: < 20; 1.07: 44 |  | 11
| 1 | 6 | 1.01: 16; 1.02: 22; 1.03: 17; 1.04: < 50; 1.05: < 20; 1.07: 34
| Mean |  | 1.01: 20; 1.02: 18.1667; 1.03: 15.5; 1.05: 23.61; 1.07: 35.6667
| Median |  | 1.01: 20.5; 1.02: 17.5; 1.03: 16; 1.05: 22; 1.07: 34
| Std Dev. |  | 1.01: 3.0332; 1.02: 2.7869; 1.03: 2.1679; 1.05: 4.1088; 1.07: 5.9889
| Rel.Std.Dev. |  | 1.01: 15.17%; 1.02: 15.34%; 1.03: 13.99%; 1.05: 17.4%; 1.07: 16.79%
| PDM3 |  | 1.01: 13.52%; 1.02: 3.115%; 1.03: -12.02%; 1.05: 34.01%; 1.07: 102.4%
| Z-Score (Absolute) |  | 1.01: 0.21; 1.02: 0.21; 1.03: 0.82; 1.04: 1.36; 1.05: 0.52; 1.07: 3.82
| Analytical results for Pr in OREAS 233b (Certified Value 1.19 ppm) | [32 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.07: 1.27 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.07: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 25g |  | 3
| 1 | 1 | 1.01: 1.257; 1.02: 1.232; 1.03: 1.1004; 1.04: 1.4; 1.05: 1.23; 1.07: 1.2 |  | 1
| 1 | 2 | 1.01: 1.243; 1.02: 1.121; 1.03: 1.0401; 1.04: 1.3; 1.05: 1.25; 1.07: 1.2 |  | 29
| 1 | 3 | 1.01: 1.265; 1.02: 1.048; 1.03: 1.0309; 1.04: 1.3; 1.05: 1.2; 1.07: 1.2 |  | 16
| 1 | 4 | 1.01: 1.272; 1.02: 1.089; 1.03: 1.05; 1.04: 1.3; 1.05: 1.26; 1.07: 1.1 |  | 1.1905
| 1 | 5 | 1.01: 1.257; 1.02: 1.1; 1.03: 1.082; 1.04: 1.3; 1.05: 1.2; 1.07: 1.1 |  | 111
| 1 | 6 | 1.01: 1.249; 1.02: 1.135; 1.03: 1.116; 1.04: 1.3; 1.05: 1.23; 1.07: 1.1
| Mean |  | 1.01: 1.2572; 1.02: 1.1208; 1.03: 1.0699; 1.04: 1.3167; 1.05: 1.2283; 1.07: 1.15
| Median |  | 1.01: 1.257; 1.02: 1.1105; 1.03: 1.066; 1.04: 1.3; 1.05: 1.23; 1.07: 1.15
| Std Dev. |  | 1.01: 0.010477; 1.02: 0.062146; 1.03: 0.03465; 1.04: 0.040825; 1.05: 0.024833; 1.07: 0.054772
| Rel.Std.Dev. |  | 1.01: 0.8334%; 1.02: 5.545%; 1.03: 3.239%; 1.04: 3.101%; 1.05: 2.022%; 1.07: 4.763%
| PDM3 |  | 1.01: 5.601%; 1.02: -5.851%; 1.03: -10.13%; 1.04: 10.6%; 1.05: 3.179%; 1.07: -3.401%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.68; 1.03: 1.18; 1.04: 1.26; 1.05: 0.39; 1.07: 0.39
| Analytical results for Pt in OREAS 233b (Indicative Value 8.33 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.11; 1.03: 1.22; 1.04: 1.27; 1.05: 1.28; 1.07: 1.34 |  | ppb
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS; 1.04: AR*MS; 1.05: AR*MS; 1.07: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 01g; 1.03: 01g; 1.04: 25g; 1.05: 0.5g; 1.07: 0.5g |  | 2
| 1 | 1 | 1.01: 9; 1.02: 9; 1.03: 6; 1.04: < 10; 1.05: 9.57; 1.07: < 10 |  | 1
| 1 | 2 | 1.01: 9; 1.02: 7; 1.03: 6; 1.04: < 10; 1.05: 9.6; 1.07: < 10 |  | 1
| 1 | 3 | 1.01: 9; 1.02: 11; 1.03: 6; 1.04: < 10; 1.05: 8.49; 1.07: < 10 |  | 16
| 1 | 4 | 1.01: 9; 1.02: 9; 1.03: 5; 1.04: < 10; 1.05: 8.59; 1.07: < 10 |  | 8.3325
| 1 | 5 | 1.01: 8; 1.02: 10; 1.03: 6; 1.04: < 10; 1.05: 9.8; 1.07: < 10 |  | 12
| 1 | 6 | 1.01: 8; 1.02: 11; 1.03: 5; 1.04: < 10; 1.05: 10.93; 1.07: < 10
| Mean |  | 1.01: 8.6667; 1.02: 9.5; 1.03: 5.6667; 1.05: 9.4967
| Median |  | 1.01: 9; 1.02: 9.5; 1.03: 6; 1.05: 9.585
| Std Dev. |  | 1.01: 0.5164; 1.02: 1.5166; 1.03: 0.5164; 1.05: 0.8942
| Rel.Std.Dev. |  | 1.01: 5.958%; 1.02: 15.96%; 1.03: 9.113%; 1.05: 9.416%
| PDM3 |  | 1.01: 4.01%; 1.02: 14.01%; 1.03: -31.99%; 1.05: 13.97%
| Z-Score (Absolute) |  | 1.01: 0.47; 1.02: 0.73; 1.03: 0.47; 1.04: 0.67; 1.05: 0.73; 1.07: 0.67
| Analytical results for Rb in OREAS 233b (Certified Value 5.88 ppm) | [32 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 5.5 | 5.85 | 5.3602 | 5.8894 | 4.6 | 6.74 | 6.92 | 6.12 | 6.5 | 5.4 | 5.6 | 5.1 | 5.6 | 5.45 | 6.3 | 5.5 | 6 | 5.8 | 5.7 | 6.3 | 5.5 | [10 empty] | 1
| 1 | 2 | 5.5 | 5.82 | 5.5958 | 5.8696 | 4.5 | 6.74 | 6.61 | 6.1 | 7.3 | 5.4 | 5.7 | 5.2 | 5.5 | 5.88 | 6.5 | 5.5 | 6.1 | 5.8 | 6 | 6.3 | 5.6 | [10 empty] | 30
| 1 | 3 | 5.5 | 5.92 | 5.6072 | 6.7737 | 4.4 | 6.65 | 6.92 | 6.05 | 6.5 | 5.5 | 5.6 | 5.4 | 5.4 | 5.91 | 6.2 | 5.5 | 6 | 5.8 | 6.2 | 6.2 | 5.5 | [10 empty] | 16
| 1 | 4 | 5.5 | 5.94 | 5.4181 | 5.8985 | 4.4 | 6.92 | 6.6 | 5.87 | 6.5 | 5.5 | 5.7 | 5.3 | 5.5 | 5.47 | 6 | 5.5 | 6.1 | 6 | 6 | 6.4 | 5.4 | [10 empty] | 5.8834
| 1 | 5 | 5.6 | 5.88 | 5.3744 | 5.7625 | 4.5 | 6.57 | 6.99 | 5.96 | 6.5 | 5.5 | 5.7 | 5.4 | 5.8 | 5.44 | 6.2 | 5 | 6.2 | 6 | 6.1 | 6.1 | 5.5 | [10 empty] | 112
| 1 | 6 | 5.6 | 5.85 | 5.5699 | 5.9275 | 4.6 | 6.88 | 6.76 | 6.11 | 6.4 | 5.4 | 5.7 | 5.5 | 5.7 | 5.67 | 6 | 5 | 6.1 | 5.7 | 5.9 | 6.2 | 5.4
| Mean |  | 5.5333 | 5.8767 | 5.4876 | 6.0202 | 4.5 | 6.75 | 6.8 | 6.035 | 6.6167 | 5.45 | 5.6667 | 5.3167 | 5.5833 | 5.6367 | 6.2 | 5.3333 | 6.0833 | 5.85 | 5.9833 | 6.25 | 5.4833
| Median |  | 5.5 | 5.865 | 5.494 | 5.894 | 4.5 | 6.74 | 6.84 | 6.075 | 6.5 | 5.45 | 5.7 | 5.35 | 5.55 | 5.57 | 6.2 | 5.5 | 6.1 | 5.8 | 6 | 6.25 | 5.5
| Std Dev. |  | 0.05164 | 0.045898 | 0.11545 | 0.37347 | 0.089443 | 0.13297 | 0.16888 | 0.10015 | 0.33714 | 0.054772 | 0.05164 | 0.1472 | 0.1472 | 0.21741 | 0.18974 | 0.2582 | 0.075277 | 0.12247 | 0.17224 | 0.10488 | 0.075277
| Rel.Std.Dev. |  | 0.9332% | 0.781% | 2.104% | 6.204% | 1.988% | 1.97% | 2.484% | 1.659% | 5.095% | 1.005% | 0.9113% | 2.769% | 2.636% | 3.857% | 3.06% | 4.841% | 1.237% | 2.094% | 2.879% | 1.678% | 1.373%
| PDM3 |  | -5.951% | -0.1151% | -6.728% | 2.325% | -23.51% | 14.73% | 15.58% | 2.576% | 12.46% | -7.367% | -3.684% | -9.633% | -5.101% | -4.194% | 5.381% | -9.35% | 3.398% | -0.5684% | 1.698% | 6.23% | -6.801%
| Z-Score (Absolute) |  | 0.61 | 0.05 | 0.7 | 0.33 | 2.6 | 1.73 | 1.83 | 0.36 | 1.48 | 0.77 | 0.35 | 1.03 | 0.51 | 0.41 | 0.67 | 1 | 0.45 | 0 | 0.26 | 0.77 | 0.71
| Analytical results for Re in OREAS 233b (Certified Value 0.002 ppm) | [32 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.007 | < 0.002 | 0.002 | 0.005 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.001 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.006 | < 0.002 | 0.003 | 0.005 | [13 empty] | 31
| 1 | 3 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.006 | < 0.002 | 0.002 | 0.006 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.002 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.007 | < 0.002 | 0.002 | 0.004 | [13 empty] | 0.0018889
| 1 | 5 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.001 | < 0.005 | 0.001 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.006 | < 0.002 | 0.003 | 0.003 | [13 empty] | 113
| 1 | 6 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.001 | < 0.005 | 0.001 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.007 | < 0.002 | 0.002 | 0.004
| Mean |  | 1.02: 0.002; 1.03: 0.002; 1.05: 0.0016667; 1.08: 0.0018333; 1.10: 0.002; 1.11: 0.0015; 1.12: 0.002; 1.16: 0.0016667; 1.18: 0.0065; 1.21: 0.0023333; 1.22: 0.0045
| Median |  | 1.02: 0.002; 1.03: 0.002; 1.05: 0.0015; 1.08: 0.002; 1.10: 0.002; 1.11: 0.0015; 1.12: 0.002; 1.16: 0.002; 1.18: 0.0065; 1.21: 0.002; 1.22: 0.0045
| Std Dev. |  | 1.02: 0; 1.03: 0; 1.05: 0.0008165; 1.08: 0.00075277; 1.10: 0; 1.11: 0.00054772; 1.12: 0; 1.16: 0.0005164; 1.18: 0.00054772; 1.21: 0.0005164; 1.22: 0.0010488
| Rel.Std.Dev. |  | 1.02: 0%; 1.03: 0%; 1.05: 48.99%; 1.08: 41.06%; 1.10: 0%; 1.11: 36.51%; 1.12: 0%; 1.16: 30.98%; 1.18: 8.427%; 1.21: 22.13%; 1.22: 23.31%
| PDM3 |  | 1.02: 5.882%; 1.03: 5.882%; 1.05: -11.76%; 1.08: -2.941%; 1.10: 5.882%; 1.11: -20.59%; 1.12: 5.882%; 1.16: -11.76%; 1.18: 244.1%; 1.21: 23.53%; 1.22: 138.2%
| Z-Score (Absolute) |  | 55.29 | 0.19 | 0.19 | 26.39 | 0.58 | 0.39 | 0.39 | 0.19 | 0.77 | 0.19 | 2887.39 | 26.39 | 26.39 | 0.58 | 5.01 | 1.35 | 0.19 | 2.7
| Analytical results for S in OREAS 233b (Certified Value 0.372 wt.%) | [32 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.38 | 0.36 | 0.37601 | 0.37419 | 0.3883 | 0.356 | 0.377 | 0.37 | 0.36 | 0.38 | 0.38 | 0.37 | 0.36 | 0.37 | 0.38 | 0.32 | 0.34302 | 0.39 | 0.32 | 0.39 | 0.35824 | 0.376 | 0.35 | 0.38 | 0.4 | 0.36 | 0.38 |  |  |  |  | 1
| 1 | 2 | 0.38 | 0.36 | 0.40364 | 0.37767 | 0.382 | 0.36 | 0.373 | 0.36 | 0.37 | 0.37 | 0.38 | 0.37 | 0.36 | 0.37 | 0.39 | 0.34 | 0.34934 | 0.34 | 0.34 | 0.38 | 0.36163 | 0.3745 | 0.34 | 0.38 | 0.39 | 0.37 | 0.39 |  |  |  |  | 16
| 1 | 3 | 0.39 | 0.36 | 0.35816 | 0.37528 | 0.44611 | 0.353 | 0.373 | 0.36 | 0.35 | 0.37 | 0.38 | 0.36 | 0.37 | 0.37 | 0.38 | 0.33 | 0.36457 | 0.37 | 0.33 | 0.38 | 0.37599 | 0.368 | 0.35 | 0.37 | 0.39 | 0.36 | 0.39 |  |  |  |  | 16
| 1 | 4 | 0.38 | 0.37 | 0.40379 | 0.36988 | 0.38901 | 0.349 | 0.368 | 0.38 | 0.36 | 0.37 | 0.37 | 0.37 | 0.37 | 0.37 | 0.38 | 0.32 | 0.34947 | 0.37 | 0.33 | 0.38 | 0.38207 | 0.375 | 0.34 | 0.39 | 0.4 | 0.36 | 0.38 |  |  |  |  | 0.37174
| 1 | 5 | 0.39 | 0.36 | 0.32607 | 0.37093 | 0.38814 | 0.351 | 0.372 | 0.37 | 0.35 | 0.38 | 0.37 | 0.37 | 0.37 | 0.37 | 0.4 | 0.33 | 0.3489 | 0.38 | 0.32 | 0.38 | 0.37928 | 0.368 | 0.35 | 0.35 | 0.39 | 0.37 | 0.39 |  |  |  |  | 114
| 1 | 6 | 0.39 | 0.36 | 0.40885 | 0.37924 | 0.38595 | 0.357 | 0.376 | 0.37 | 0.35 | 0.38 | 0.37 | 0.36 | 0.37 | 0.38 | 0.39 | 0.33 | 0.35324 | 0.34 | 0.34 | 0.38 | 0.38307 | 0.364 | 0.35 | 0.37 | 0.39 | 0.36 | 0.38
| Mean |  | 0.385 | 0.36167 | 0.37942 | 0.37453 | 0.39658 | 0.35433 | 0.37317 | 0.36833 | 0.35667 | 0.375 | 0.375 | 0.36667 | 0.36667 | 0.37167 | 0.38667 | 0.32833 | 0.35142 | 0.365 | 0.33 | 0.38167 | 0.37338 | 0.37092 | 0.34667 | 0.37333 | 0.39333 | 0.36333 | 0.385
| Median |  | 0.385 | 0.36 | 0.38982 | 0.37474 | 0.38822 | 0.3545 | 0.373 | 0.37 | 0.355 | 0.375 | 0.375 | 0.37 | 0.37 | 0.37 | 0.385 | 0.33 | 0.3494 | 0.37 | 0.33 | 0.38 | 0.37764 | 0.37125 | 0.35 | 0.375 | 0.39 | 0.36 | 0.385
| Std Dev. |  | 0.0054772 | 0.0040825 | 0.032732 | 0.0036696 | 0.024397 | 0.0040825 | 0.0031885 | 0.0075277 | 0.008165 | 0.0054772 | 0.0054772 | 0.005164 | 0.005164 | 0.0040825 | 0.008165 | 0.0075277 | 0.0072269 | 0.020736 | 0.0089443 | 0.0040825 | 0.010755 | 0.0049032 | 0.005164 | 0.013663 | 0.005164 | 0.005164 | 0.0054772
| Rel.Std.Dev. |  | 1.423% | 1.129% | 8.627% | 0.9798% | 6.152% | 1.152% | 0.8544% | 2.044% | 2.289% | 1.461% | 1.461% | 1.408% | 1.408% | 1.098% | 2.112% | 2.293% | 2.056% | 5.681% | 2.71% | 1.07% | 2.88% | 1.322% | 1.49% | 3.66% | 1.313% | 1.421% | 1.423%
| PDM3 |  | 3.566% | -2.71% | 2.065% | 0.7508% | 6.682% | -4.683% | 0.3832% | -0.917% | -4.055% | 0.8763% | 0.8763% | -1.365% | -1.365% | -0.02035% | 4.015% | -11.68% | -5.466% | -1.814% | -11.23% | 2.67% | 0.4408% | -0.2221% | -6.745% | 0.428% | 5.808% | -2.262% | 3.566%
| Z-Score (Absolute) |  | 1.08 | 0.81 | 0.63 | 0.23 | 2.02 | 1.4 | 0.12 | 0.27 | 1.21 | 0.27 | 0.27 | 0.4 | 0.4 | 0 | 1.21 | 3.51 | 1.64 | 0.54 | 3.37 | 0.81 | 0.14 | 0.06 | 2.02 | 0.13 | 1.75 | 0.67 | 1.08
| Analytical results for Sb in OREAS 233b (Certified Value 0.79 ppm) | [32 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 1.04 | 0.94 | 2 | 1.0285 | 0.82 | 1 | 0.25 | 0.41 | < 5 | 0.94 | 0.82 | 0.84 | 0.89 | 74.987 | 0.82 | < 5 | < 5 | 0.78 | 0.6 | 0.58 | 5.91 | 0.5 | 1.0601 | 0.91 | 0.219 | 0.59 | 0.45 | 0.52 | 0.64 |  |  | 1
| 1 | 2 | 1.04 | 0.9 | 4.63 | 0.92687 | 0.79 | 1.03 | 0.27 | 0.44 | < 5 | 0.98 | 0.81 | 0.9 | 0.92 | 76.429 | 0.83 | < 5 | < 5 | 0.86 | 0.68 | 0.58 | 6.05 | 0.5 | 1.1329 | 0.9 | 0.209 | 0.62 | 0.48 | 0.53 | 0.64 |  |  | 32
| 1 | 3 | 1.05 | 0.92 | 2.74 | 1.0983 | 0.73 | 0.99 | 0.28 | 0.39 | < 5 | 0.99 | 0.81 | 0.86 | 0.92 | 78.839 | 0.79 | < 5 | < 5 | 0.88 | 0.63 | 0.61 | 8.3 | 0.5 | 1.0484 | 0.86 | 0.22 | 0.61 | 0.5 | 0.49 | 0.62 |  |  | 16
| 1 | 4 | 1.04 | 0.9 | 5.38 | 0.96484 | 0.72 | 1.01 | 0.27 | 0.42 | < 5 | 1 | 0.83 | 0.88 | 0.92 | 78.322 | 0.8 | < 5 | < 5 | 0.8 | 0.68 | 0.55 | 7.83 | 0.5 | 1.0318 | 0.89 | 0.192 | 0.57 | 0.47 | 0.51 | 0.61 |  |  | 0.78655
| 1 | 5 | 1.08 | 0.92 | 3.37 | 0.90569 | 0.76 | 0.95 | 0.25 | 0.43 | < 5 | 1.06 | 0.78 | 0.86 | 0.93 | 78.025 | 0.81 | < 5 | < 5 | 0.87 | 0.63 | 0.55 | 5.81 | 0.5 | 0.97222 | 0.95 | 0.2 | 0.54 | 0.48 | 0.49 | 0.61 |  |  | 115
| 1 | 6 | 1.05 | 0.91 | 4.57 | 0.9234 | 0.78 | 1 | 0.22 | 0.42 | < 5 | 1.14 | 0.79 | 0.91 | 0.94 | 84.562 | 0.79 | < 5 | < 5 | 0.93 | 0.67 | 0.54 | 8.76 | 0.5 | 1.0193 | 0.88 | 0.221 | 0.53 | 0.47 | 0.51 | 0.58
| Mean |  | 1.05 | 0.915 | 3.7817 | 0.97459 | 0.76667 | 0.99667 | 0.25667 | 0.41833 |  | 1.0183 | 0.80667 | 0.875 | 0.92 | 78.5273 | 0.80667 |  |  | 0.85333 | 0.64833 | 0.56833 | 7.11 | 0.5 | 1.0441 | 0.89833 | 0.21017 | 0.57667 | 0.475 | 0.50833 | 0.61667
| Median |  | 1.045 | 0.915 | 3.97 | 0.94585 | 0.77 | 1 | 0.26 | 0.42 |  | 0.995 | 0.81 | 0.87 | 0.92 | 78.1735 | 0.805 |  |  | 0.865 | 0.65 | 0.565 | 6.94 | 0.5 | 1.0401 | 0.895 | 0.214 | 0.58 | 0.475 | 0.51 | 0.615
| Std Dev. |  | 0.015492 | 0.015166 | 1.2903 | 0.074757 | 0.037771 | 0.026583 | 0.021602 | 0.017224 |  | 0.07111 | 0.018619 | 0.026646 | 0.016733 | 3.2786 | 0.01633 |  |  | 0.055015 | 0.033116 | 0.026394 | 1.335 | 0 | 0.053091 | 0.030605 | 0.012057 | 0.036697 | 0.016432 | 0.016021 | 0.022509
| Rel.Std.Dev. |  | 1.475% | 1.657% | 34.12% | 7.671% | 4.927% | 2.667% | 8.417% | 4.117% |  | 6.983% | 2.308% | 3.045% | 1.819% | 4.175% | 2.024% |  |  | 6.447% | 5.108% | 4.644% | 18.78% | 0% | 5.085% | 3.407% | 5.737% | 6.364% | 3.459% | 3.152% | 3.65%
| PDM3 |  | 33.49% | 16.33% | 380.8% | 23.91% | -2.528% | 26.71% | -67.37% | -46.81% |  | 29.47% | 2.557% | 11.24% | 16.97% | 9884% | 2.557% |  |  | 8.49% | -17.57% | -27.74% | 803.9% | -36.43% | 32.75% | 14.21% | -73.28% | -26.68% | -39.61% | -35.37% | -21.6%
| Z-Score (Absolute) |  | 0.43 | 0.08 | 7.68 | 0.23 | 0.32 | 0.29 | 1.67 | 1.24 | 4.28 | 0.35 | 0.21 | 0.03 | 0.09 | 205.98 | 0.21 | 4.28 | 4.28 | 0.09 | 0.63 | 0.84 | 16.51 | NA | 0.42 | 0.03 | 1.79 | 0.82 | 1.09 | 1 | 0.72
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 233b (Certified Value 4.78 ppm) | [32 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 4 | 4.9 | 5.365 | 6.02 | 5.042 | 6.7 | 7.4 | 6 | 5.7 | 5 | 4.7 | 5 | 4.5 | 4.5 | 4.5 | 4 | 6 | 3.9 | 5 | 4 | 4.9 | 1.5 | 4.9 | 4.7 | 3.6 | [6 empty] | 1
| 1 | 2 | 4.2 | 4.89 | 4.947 | 6.31 | 5.0163 | 6.4 | 7.4 | 6 | 5.7 | 5 | 4.8 | 5 | 4.5 | 4.8 | 4.6 | 4 | 6 | 4.2 | 4.8 | 4 | 4.9 | 1.9 | 4.8 | 5 | 3.8 | [6 empty] | 33
| 1 | 3 | 4.3 | 5 | 4.83 | 6.54 | 5.3698 | 6.2 | 6.8 | 6 | 5.8 | 5 | 4.5 | 5.1 | 4.4 | 4.8 | 4.6 | 4 | 6 | 4 | 4.7 | 3 | 4.8 | 1.7 | 4.8 | 4.7 | 3.7 | [6 empty] | 16
| 1 | 4 | 4.4 | 4.99 | 4.651 | 6.06 | 4.9393 | 5.6 | 6.9 | 6 | 5.7 | 5 | 4.7 | 5 | 4.5 | 4.9 | 4.6 | 4 | 5 | 3.8 | 4.6 | 3 | 4.9 | 1.6 | 4.8 | 4.8 | 3.6 | [6 empty] | 4.7753
| 1 | 5 | 4.3 | 4.91 | 4.671 | 6.3 | 4.8854 | 6.1 | 6.7 | 6 | 5.6 | 5 | 4.4 | 5.1 | 4.5 | 4.8 | 4.7 | 4 | 5 | 4 | 4.8 | 3 | 4.9 | 1.7 | 4.9 | 4.6 | 3.7 | [6 empty] | 116
| 1 | 6 | 4.2 | 4.91 | 5.596 | 6.43 | 5.0487 | 6.4 | 7 | 6 | 5.9 | 5 | 4.5 | 5 | 4.5 | 5 | 4.8 | 4 | 5 | 4 | 4.5 | 3 | 4.8 | 1.4 | 5 | 4.6 | 3.7
| Mean |  | 4.2333 | 4.9333 | 5.01 | 6.2767 | 5.0502 | 6.2333 | 7.0333 | 6 | 5.7333 | 5 | 4.6 | 5.0333 | 4.4833 | 4.8 | 4.6333 | 4 | 5.5 | 3.9833 | 4.7333 | 3.3333 | 4.8667 | 1.6333 | 4.8667 | 4.7333 | 3.6833
| Median |  | 4.25 | 4.91 | 4.8885 | 6.305 | 5.0291 | 6.3 | 6.95 | 6 | 5.7 | 5 | 4.6 | 5 | 4.5 | 4.8 | 4.6 | 4 | 5.5 | 4 | 4.75 | 3 | 4.9 | 1.65 | 4.85 | 4.7 | 3.7
| Std Dev. |  | 0.13663 | 0.048442 | 0.38719 | 0.20363 | 0.169 | 0.37238 | 0.30111 | 0 | 0.10328 | 0 | 0.15492 | 0.05164 | 0.040825 | 0.16733 | 0.10328 | 0 | 0.54772 | 0.13292 | 0.17512 | 0.5164 | 0.05164 | 0.17512 | 0.08165 | 0.15055 | 0.075277
| Rel.Std.Dev. |  | 3.227% | 0.9819% | 7.728% | 3.244% | 3.346% | 5.974% | 4.281% | 0% | 1.801% | 0% | 3.368% | 1.026% | 0.9106% | 3.486% | 2.229% | 0% | 9.959% | 3.337% | 3.7% | 15.49% | 1.061% | 10.72% | 1.678% | 3.181% | 2.044%
| PDM3 |  | -11.35% | 3.309% | 4.915% | 31.44% | 5.757% | 30.53% | 47.29% | 25.65% | 20.06% | 4.705% | -3.671% | 5.403% | -6.114% | 0.517% | -2.973% | -16.24% | 15.18% | -16.58% | -0.879% | -30.2% | 1.913% | -65.8% | 1.913% | -0.879% | -22.87%
| Z-Score (Absolute) |  | 1.8 | 0.3 | 0.53 | 4.32 | 0.65 | 4.19 | 6.59 | NA | 2.7 | NA | 0.7 | 0.6 | 1.05 | 0.1 | 0.6 | NA | NA | 2.55 | 0.3 | NA | 0.1 | 9.59 | 0.1 | 0.3 | 3.44
Results from laboratories 1.10, 1.12, 1.19, 1.22 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 233b (Indicative Value 0.59 ppm) | [32 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | 0.42 | < 1 | < 1.5 | 1.1 | < 1 | < 1 | < 1 | < 5 | 2 | 0.5 | 0.5 | 0.7 | 0.9 | < 5 | < 5 | < 1 | 0.5 | 1.52 | < 1 | 0.47897 | 0.9 | 0.605 | 1.6 | < 1 | 0.5 | 0.5 |  |  |  |  | 1
| 1 | 2 | < 1 | 0.42 | < 1 | < 1.5 | 0.8 | 2 | < 1 | < 1 | < 5 | 2.2 | 0.6 | 0.6 | 0.9 | 1.1 | < 5 | < 5 | < 1 | 0.5 | 1.46 | < 1 | 0.36366 | 0.6 | 0.531 | 2.8 | < 1 | 0.4 | 0.4 |  |  |  |  | 7
| 1 | 3 | < 1 | 0.47 | < 1 | < 1.5 | 0.8 | 2 | < 1 | < 1 | < 5 | 1.7 | 0.8 | 0.6 | 0.5 | 0.8 | < 5 | < 5 | < 1 | 0.6 | 1.95 | < 1 | 0.38505 | 0.8 | 0.527 | 2.2 | < 1 | 0.4 | 0.5 |  |  |  |  | 16
| 1 | 4 | < 1 | 0.44 | < 1 | < 1.5 | 0.6 | 2 | < 1 | < 1 | < 5 | 2.2 | 0.8 | 0.6 | 0.5 | 1 | < 5 | < 5 | < 1 | 0.5 | 3.14 | < 1 | 0.4365 | 0.6 | 0.413 | 0.9 | < 1 | 0.4 | 0.5 |  |  |  |  | 0.59182
| 1 | 5 | < 1 | 0.39 | < 1 | < 1.5 | 0.6 | 4 | < 1 | < 1 | < 5 | 2.4 | 0.9 | 0.6 | 0.6 | 1.1 | < 5 | < 5 | < 1 | 0.5 | 4.06 | < 1 | 0.39248 | 0.7 | 0.416 | 0.4 | < 1 | 0.5 | 0.5 |  |  |  |  | 13
| 1 | 6 | < 1 | 0.39 | < 1 | < 1.5 | 0.7 | < 1 | < 1 | < 1 | < 5 | 2.3 | 0.7 | 0.6 | 0.8 | 0.8 | < 5 | < 5 | < 1 | 0.5 | 7.39 | < 1 | 0.393 | 0.6 | 0.424 | 1.9 | < 1 | 0.4 | 0.4
| Mean |  | 1.02: 0.42167; 1.07: 0.76667; 1.08: 2.5; 1.13: 2.1333; 1.14: 0.71667; 1.15: 0.58333; 1.16: 0.66667; 1.18: 0.95; 1.23: 0.51667; 1.25: 3.2533; 1.28: 0.40828; 1.29: 0.7; 1.30: 0.486; 1.31: 1.6333; 1.33: 0.43333; 1.34: 0.46667
| Median |  | 1.02: 0.42; 1.07: 0.75; 1.08: 2; 1.13: 2.2; 1.14: 0.75; 1.15: 0.6; 1.16: 0.65; 1.18: 0.95; 1.23: 0.5; 1.25: 2.545; 1.28: 0.39274; 1.29: 0.65; 1.30: 0.4755; 1.31: 1.75; 1.33: 0.4; 1.34: 0.5
| Std Dev. |  | 1.02: 0.030605; 1.07: 0.18619; 1.08: 1; 1.13: 0.25033; 1.14: 0.1472; 1.15: 0.040825; 1.16: 0.1633; 1.18: 0.13784; 1.23: 0.040825; 1.25: 2.2669; 1.28: 0.041969; 1.29: 0.12649; 1.30: 0.079925; 1.31: 0.87331; 1.33: 0.05164; 1.34: 0.05164
| Rel.Std.Dev. |  | 1.02: 7.258%; 1.07: 24.29%; 1.08: 40%; 1.13: 11.73%; 1.14: 20.54%; 1.15: 6.999%; 1.16: 24.49%; 1.18: 14.51%; 1.23: 7.902%; 1.25: 69.68%; 1.28: 10.28%; 1.29: 18.07%; 1.30: 16.45%; 1.31: 53.47%; 1.33: 11.92%; 1.34: 11.07%
| PDM3 |  | 1.02: -28.75%; 1.07: 29.54%; 1.08: 322.4%; 1.13: 260.5%; 1.14: 21.1%; 1.15: -1.434%; 1.16: 12.65%; 1.18: 60.52%; 1.23: -12.7%; 1.25: 449.7%; 1.28: -31.01%; 1.29: 18.28%; 1.30: -17.88%; 1.31: 176%; 1.33: -26.78%; 1.34: -21.15%
| Z-Score (Absolute) |  | 0.42 | 0.82 | 0.42 | 0.84 | 0.93 | 6.32 | 0.42 | 0.42 | 9.69 | 7.84 | 0.67 | 0 | 0.42 | 1.85 | 9.69 | 9.69 | 0.42 | 0.34 | 13.5 | 0.42 | 0.89 | 0.59 | 0.49 | 5.31 | 0.42 | 0.76 | 0.59
| Analytical results for Si in OREAS 233b (Indicative Value 0.071 wt.%) | [32 empty] | Indicative
Si | Round | Replicate | 1.01: Lab; 1.02: Lab |  |  |  |  | 1
| No | No | 1.01: 1.25; 1.02: 1.27 |  |  |  |  | wt.%
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS |  |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 25g |  |  |  |  | 3
| 1 | 1 | 1.01: 0.071154; 1.02: 0.0665 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.066508; 1.02: 0.088 |  |  |  |  | 5
| 1 | 3 | 1.01: 0.063382; 1.02: 0.07 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.071686; 1.02: 0.066 |  |  |  |  | 0.071459
| 1 | 5 | 1.01: 0.077737; 1.02: 0.0825 |  |  |  |  | 11
| 1 | 6 | 1.01: 0.058043; 1.02: 0.076
| Mean |  | 1.01: 0.068085; 1.02: 0.074833
| Median |  | 1.01: 0.068831; 1.02: 0.073
| Std Dev. |  | 1.01: 0.0069425; 1.02: 0.0089926
| Rel.Std.Dev. |  | 1.01: 10.2%; 1.02: 12.02%
| PDM3 |  | 1.01: -4.722%; 1.02: 4.722%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Sm in OREAS 233b (Certified Value 1.51 ppm) | [32 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.05: 1.07; 1.07: 1.08; 1.08: 1.27 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.05: 0.5g; 1.07: 0.5g; 1.08: 25g |  |  |  |  | 2
| 1 | 1 | 1.01: 1.542; 1.02: 1.509; 1.03: 1.3712; 1.05: 1.5; 1.07: 1.68; 1.08: 1.5 |  |  |  |  | 1
| 1 | 2 | 1.01: 1.502; 1.02: 1.451; 1.03: 1.411; 1.05: 1.7; 1.07: 1.67; 1.08: 1.4 |  |  |  |  | 18
| 1 | 3 | 1.01: 1.515; 1.02: 1.328; 1.03: 1.4085; 1.05: 1.9; 1.07: 1.65; 1.08: 1.3 |  |  |  |  | 16
| 1 | 4 | 1.01: 1.541; 1.02: 1.361; 1.03: 1.3168; 1.05: 1.7; 1.07: 1.7; 1.08: 1.3 |  |  |  |  | 1.5135
| 1 | 5 | 1.01: 1.541; 1.02: 1.422; 1.03: 1.3403; 1.05: 1.7; 1.07: 1.67; 1.08: 1.4 |  |  |  |  | 117
| 1 | 6 | 1.01: 1.55; 1.02: 1.457; 1.03: 1.4085; 1.05: 1.8; 1.07: 1.74; 1.08: 1.2
| Mean |  | 1.01: 1.5318; 1.02: 1.4213; 1.03: 1.3761; 1.05: 1.7167; 1.07: 1.685; 1.08: 1.35
| Median |  | 1.01: 1.541; 1.02: 1.4365; 1.03: 1.3898; 1.05: 1.7; 1.07: 1.675; 1.08: 1.35
| Std Dev. |  | 1.01: 0.018841; 1.02: 0.066617; 1.03: 0.040337; 1.05: 0.13292; 1.07: 0.031464; 1.08: 0.10488
| Rel.Std.Dev. |  | 1.01: 1.23%; 1.02: 4.687%; 1.03: 2.931%; 1.05: 7.743%; 1.07: 1.867%; 1.08: 7.769%
| PDM3 |  | 1.01: 1.213%; 1.02: -6.088%; 1.03: -9.08%; 1.05: 13.43%; 1.07: 11.33%; 1.08: -10.8%
| Z-Score (Absolute) |  | 1.01: 0.33; 1.02: 0.33; 1.03: 0.6; 1.05: 1.43; 1.07: 1.24; 1.08: 0.75
| Analytical results for Sn in OREAS 233b (Certified Value 0.63 ppm) | [32 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.6 | 0.62 | < 1 | 0.57009 | 0.56445 | 0.72 | 0.59 | 0.8 | 0.7 | < 10 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | < 10 | < 5 | 0.7 | 0.6 | 41.7 | < 0.5 | 0.69554 | 0.6 | 1.2 | 0.6 | 0.7 | 0.4 |  |  |  |  | 1
| 1 | 2 | 0.7 | 0.62 | < 1 | 0.57222 | 0.63016 | 0.69 | 0.58 | 0.8 | 0.69 | < 10 | 0.7 | 0.6 | 0.7 | 0.6 | 0.6 | < 10 | < 5 | 0.6 | 0.6 | 40.39 | < 0.5 | 0.68337 | 0.6 | 1.1 | 0.7 | 0.8 | 0.4 |  |  |  |  | 19
| 1 | 3 | 0.6 | 0.64 | < 1 | 0.57174 | 0.61581 | 0.66 | 0.58 | 0.9 | 0.68 | < 10 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | < 10 | < 5 | 0.7 | 0.6 | 43.82 | < 0.5 | 0.63542 | 0.6 | 1.2 | 0.7 | 0.7 | 0.4 |  |  |  |  | 16
| 1 | 4 | 0.7 | 0.65 | < 1 | 0.54507 | 0.61525 | 0.66 | 0.56 | 0.9 | 0.67 | < 10 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | < 10 | < 5 | 0.6 | 0.6 | 39.51 | < 0.5 | 0.62155 | 0.6 | 1.1 | 0.7 | 0.7 | 0.4 |  |  |  |  | 0.62857
| 1 | 5 | 0.7 | 0.63 | < 1 | 0.59403 | 0.58498 | 0.7 | 0.56 | 0.9 | 0.64 | < 10 | 0.5 | 0.6 | 0.6 | 0.6 | 0.6 | < 10 | < 5 | 0.6 | 0.6 | 41.03 | < 0.5 | 0.64636 | 0.6 | 1 | 0.6 | 0.7 | 0.4 |  |  |  |  | 118
| 1 | 6 | 0.7 | 0.61 | < 1 | 0.58358 | 0.53084 | 0.72 | 0.58 | 0.9 | 0.69 | < 10 | 0.6 | 0.6 | 0.6 | 0.7 | 0.6 | < 10 | < 5 | 0.6 | 0.6 | 40.6 | < 0.5 | 0.78504 | 0.6 | 0.9 | 0.7 | 0.7 | 0.4
| Mean |  | 0.66667 | 0.62833 |  | 0.57279 | 0.59025 | 0.69167 | 0.575 | 0.86667 | 0.67833 |  | 0.6 | 0.6 | 0.61667 | 0.61667 | 0.6 |  |  | 0.63333 | 0.6 | 41.175 |  | 0.67788 | 0.6 | 1.0833 | 0.66667 | 0.71667 | 0.4
| Median |  | 0.7 | 0.625 |  | 0.57198 | 0.60012 | 0.695 | 0.58 | 0.9 | 0.685 |  | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 |  |  | 0.6 | 0.6 | 40.815 |  | 0.66487 | 0.6 | 1.1 | 0.7 | 0.7 | 0.4
| Std Dev. |  | 0.05164 | 0.01472 |  | 0.016399 | 0.03767 | 0.027142 | 0.012247 | 0.05164 | 0.02137 |  | 0.063246 | 0 | 0.040825 | 0.040825 | 0 |  |  | 0.05164 | 0 | 1.484 |  | 0.059657 | 0 | 0.1169 | 0.05164 | 0.040825 | 6.0809e-17
| Rel.Std.Dev. |  | 7.746% | 2.343% |  | 2.863% | 6.382% | 3.924% | 2.13% | 5.958% | 3.15% |  | 10.54% | 0% | 6.62% | 6.62% | 0% |  |  | 8.154% | 0% | 3.604% |  | 8.801% | 0% | 10.79% | 7.746% | 5.696% | 1.52e-14%
| PDM3 |  | 6.061% | -0.03756% |  | -8.874% | -6.097% | 10.04% | -8.522% | 37.88% | 7.917% |  | -4.545% | -4.545% | -1.894% | -1.894% | -4.545% |  |  | 0.7579% | -4.545% | 6451% |  | 7.845% | -4.545% | 72.35% | 6.061% | 14.02% | -36.36%
| Z-Score (Absolute) |  | 0.52 | 0 | 1.73 | 0.75 | 0.51 | 0.85 | 0.72 | 3.21 | 0.67 | 58.96 | 0.38 | 0.38 | 0.16 | 0.16 | 0.38 | 58.96 | 25.24 | 0.07 | 0.38 | 546.82 | 5.1 | 0.67 | 0.38 | 6.14 | 0.52 | 1.19 | 3.08
| Analytical results for Sr in OREAS 233b (Certified Value 36.6 ppm) | [32 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 22.2 | 36.5 | 155.25 | 49.2639 | 38.7903 | 52.7 | 46 | 46.4 | 47.05 | 42 | 38 | 39.1 | 37.2 | 33.2 | 34.9 | 33 | 40.1761 | 37.9 | 31.4 | 37.3 | 28.5 | 34.6518 | 26 | 36.1 | 30.3 | 32.5 | 36.9 | 27.5 |  |  |  | 1
| 1 | 2 | 22.1 | 36.21 | 153.67 | 49.4166 | 38.3643 | 54.5 | 46 | 46.8 | 45.41 | 41 | 37 | 39 | 38.4 | 34.4 | 36.2 | 37 | 43.9108 | 38.3 | 33.1 | 37.1 | 25.82 | 35.5539 | 27 | 37 | 33 | 33.3 | 38.4 | 27.7 |  |  |  | 20
| 1 | 3 | 21.7 | 37.44 | 160.56 | 48.9692 | 41.3226 | 50.8 | 47 | 49.1 | 47.9 | 40 | 37 | 40 | 38.4 | 34.1 | 34.4 | 36 | 46.5509 | 38.4 | 31.8 | 36.7 | 30.34 | 37.9068 | 25 | 36.7 | 31.7 | 33.9 | 37.6 | 27.2 |  |  |  | 16
| 1 | 4 | 28.6 | 37.3 | 146.43 | 49.3046 | 37.207 | 51.6 | 42 | 49.5 | 46.77 | 42 | 37 | 37.8 | 37.3 | 34.8 | 35.4 | 35 | 46.5618 | 35.7 | 30.2 | 35.9 | 28.68 | 38.9763 | 26 | 36.2 | 31.1 | 34.1 | 37.2 | 26.5 |  |  |  | 36.6157
| 1 | 5 | 23 | 36.5 | 146.97 | 49.8564 | 36.7754 | 55 | 43 | 51.2 | 43.95 | 42 | 35 | 39.6 | 38.2 | 34.4 | 36.3 | 34 | 43.9066 | 34.8 | 30.8 | 35.2 | 27.44 | 38.2788 | 24 | 36.2 | 30.1 | 34.1 | 36.1 | 27.4 |  |  |  | 119
| 1 | 6 | 24.4 | 36.29 | 165.59 | 50.0332 | 38.4592 | 52.9 | 47 | 48.8 | 48.29 | 42 | 35 | 38.4 | 38.3 | 35.9 | 36.6 | 33 | 44.939 | 36.3 | 31.3 | 34.7 | 29.82 | 37.3767 | 24 | 35.9 | 28.6 | 32.6 | 36.5 | 26.8
| Mean |  | 23.6667 | 36.7067 | 154.745 | 49.474 | 38.4865 | 52.9167 | 45.1667 | 48.6333 | 46.5617 | 41.5 | 36.5 | 38.9833 | 37.9667 | 34.4667 | 35.6333 | 34.6667 | 44.3409 | 36.9 | 31.4333 | 36.15 | 28.4333 | 37.124 | 25.3333 | 36.35 | 30.8 | 33.4167 | 37.1167 | 27.1833
| Median |  | 22.6 | 36.5 | 154.46 | 49.3606 | 38.4118 | 52.8 | 46 | 48.95 | 46.91 | 42 | 37 | 39.05 | 38.25 | 34.4 | 35.8 | 34.5 | 44.4249 | 37.1 | 31.35 | 36.3 | 28.59 | 37.6418 | 25.5 | 36.2 | 30.7 | 33.6 | 37.05 | 27.3
| Std Dev. |  | 2.5997 | 0.52831 | 7.5157 | 0.39748 | 1.5948 | 1.6192 | 2.137 | 1.7851 | 1.6253 | 0.83666 | 1.2247 | 0.79603 | 0.56095 | 0.88468 | 0.87331 | 1.633 | 2.3619 | 1.5113 | 0.98522 | 1.0578 | 1.6395 | 1.6744 | 1.2111 | 0.41352 | 1.5047 | 0.73326 | 0.81833 | 0.4535
| Rel.Std.Dev. |  | 10.98% | 1.439% | 4.857% | 0.8034% | 4.144% | 3.06% | 4.731% | 3.671% | 3.491% | 2.016% | 3.355% | 2.042% | 1.477% | 2.567% | 2.451% | 4.711% | 5.327% | 4.096% | 3.134% | 2.926% | 5.766% | 4.51% | 4.781% | 1.138% | 4.885% | 2.194% | 2.205% | 1.668%
| PDM3 |  | -35.36% | 0.2486% | 322.6% | 35.12% | 5.109% | 44.52% | 23.35% | 32.82% | 27.16% | 13.34% | -0.3159% | 6.466% | 3.69% | -5.869% | -2.683% | -5.323% | 21.1% | 0.7766% | -14.15% | -1.272% | -22.35% | 1.388% | -30.81% | -0.7255% | -15.88% | -8.737% | 1.368% | -25.76%
| Z-Score (Absolute) |  | 2.19 | 0.02 | 19.68 | 2.11 | 0.28 | 2.69 | 1.4 | 1.97 | 1.63 | 0.78 | 0.05 | 0.36 | 0.19 | 0.39 | 0.2 | 0.36 | 1.26 | 0.02 | 0.9 | 0.11 | 1.4 | 0.05 | 1.91 | 0.08 | 1 | 0.57 | 0.05 | 1.6
| Analytical results for Ta in OREAS 233b (Certified Value < 0.01 ppm) | [32 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.22 | 0.08 | < 0.01 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.15 | 0.16 | < 0.01 | < 0.01 | [11 empty] | 21
| 1 | 3 | < 0.05 | < 0.01 | 0.011187 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 1.15 | 0.1 | < 0.01 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.33 | 0.13 | < 0.01 | < 0.01 | [11 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.54 | 0.16 | < 0.01 | < 0.01 | [11 empty] | 120
| 1 | 6 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.95 | 0.09 | < 0.01 | < 0.01
| Mean |  | 1.03: 0.011187; 1.05: 0.03; 1.10: 5; 1.12: 0.01; 1.13: 0.01; 1.19: 0.01; 1.21: 0.55667; 1.22: 0.12
| Median |  | 1.03: 0.011187; 1.05: 0.03; 1.10: 5; 1.12: 0.01; 1.13: 0.01; 1.19: 0.01; 1.21: 0.435; 1.22: 0.115
| Std Dev. |  | 1.05: 0; 1.10: 0; 1.13: 0; 1.21: 0.40918; 1.22: 0.035214
| Rel.Std.Dev. |  | 1.05: 0%; 1.10: 0%; 1.13: 0%; 1.21: 73.51%; 1.22: 29.34%
PDM3
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.62 | 0.4 | 0.67 | 0.4 | 0.67 | 268.78 | 1.75 | 0.63 | 0.4 | 0.67 | 0.67 | 0.4 | 12.54 | 0.63 | 29.09 | 5.53 | 0.67 | 0.67
| Analytical results for Tb in OREAS 233b (Certified Value 0.3 ppm) | [32 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.04; 1.05: 1.07; 1.07: 1.08; 1.08: 1.27; 1.10: 1.32 |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.10: AR*OES/MS |  |  |  | 2
|  |  | 1.01: 0.25g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.07: 0.5g; 1.08: 25g; 1.10: 0.25g |  |  |  | 3
| 1 | 1 | 1.01: 0.29; 1.02: 0.303; 1.03: 0.313; 1.04: 0.30926; 1.05: 0.4; 1.07: 0.31; 1.08: 0.25; 1.10: 0.31 |  |  |  | 1
| 1 | 2 | 1.01: 0.29; 1.02: 0.303; 1.03: 0.304; 1.04: 0.30821; 1.05: 0.4; 1.07: 0.31; 1.08: 0.25; 1.10: 0.33 |  |  |  | 22
| 1 | 3 | 1.01: 0.29; 1.02: 0.31; 1.03: 0.287; 1.04: 0.31247; 1.05: 0.3; 1.07: 0.3; 1.08: 0.25; 1.10: 0.32 |  |  |  | 16
| 1 | 4 | 1.01: 0.3; 1.02: 0.306; 1.03: 0.282; 1.04: 0.29639; 1.05: 0.3; 1.07: 0.3; 1.08: 0.25; 1.10: 0.32 |  |  |  | 0.30405
| 1 | 5 | 1.01: 0.29; 1.02: 0.304; 1.03: 0.292; 1.04: 0.3042; 1.05: 0.4; 1.07: 0.3; 1.08: 0.25; 1.10: 0.32 |  |  |  | 121
| 1 | 6 | 1.01: 0.29; 1.02: 0.307; 1.03: 0.311; 1.04: 0.31342; 1.05: 0.4; 1.07: 0.31; 1.08: 0.25; 1.10: 0.3
| Mean |  | 1.01: 0.29167; 1.02: 0.3055; 1.03: 0.29817; 1.04: 0.30733; 1.05: 0.36667; 1.07: 0.305; 1.08: 0.25; 1.10: 0.31667
| Median |  | 1.01: 0.29; 1.02: 0.305; 1.03: 0.298; 1.04: 0.30874; 1.05: 0.4; 1.07: 0.305; 1.08: 0.25; 1.10: 0.32
| Std Dev. |  | 1.01: 0.0040825; 1.02: 0.0027386; 1.03: 0.012983; 1.04: 0.0062852; 1.05: 0.05164; 1.07: 0.0054772; 1.08: 0; 1.10: 0.010328
| Rel.Std.Dev. |  | 1.01: 1.4%; 1.02: 0.8964%; 1.03: 4.354%; 1.04: 2.045%; 1.05: 14.08%; 1.07: 1.796%; 1.08: 0%; 1.10: 3.261%
| PDM3 |  | 1.01: -4.074%; 1.02: 0.4754%; 1.03: -1.936%; 1.04: 1.077%; 1.05: 20.59%; 1.07: 0.3109%; 1.08: -17.78%; 1.10: 4.148%
| Z-Score (Absolute) |  | 1.01: 0.99; 1.02: 0.02; 1.03: 0.52; 1.04: 0.15; 1.05: 4.48; 1.07: 0.02; 1.08: 4.03; 1.10: 0.83
| Analytical results for Te in OREAS 233b (Certified Value 0.089 ppm) | [32 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.08 | 0.09 | 0.12349 | 0.1 | 0.12 | < 0.05 | < 0.1 | < 5 | 0.1 | 0.09 | 0.07 | 0.07 | 0.07 | < 5 | 0.08 | 0.09 | < 0.1 | 0.13251 | 0.11 | 0.54 | 0.04 | 0.09 | 0.08 | 0.07 | [7 empty] | 1
| 1 | 2 | 0.06 | 0.09 | 0.10833 | 0.07 | 0.09 | < 0.05 | 0.1 | < 5 | < 0.1 | 0.1 | 0.08 | 0.09 | 0.1 | < 5 | 0.1 | 0.09 | < 0.1 | 0.15706 | 0.1 | 0.437 | 0.04 | 0.1 | 0.09 | 0.12 | [7 empty] | 23
| 1 | 3 | 0.1 | 0.09 | 0.12077 | 0.08 | 0.08 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.09 | 0.08 | 0.09 | 0.11 | < 5 | 0.1 | 0.08 | < 0.1 | 0.06812 | 0.09 | 0.334 | 0.03 | 0.09 | 0.09 | 0.07 | [7 empty] | 16
| 1 | 4 | < 0.05 | 0.09 | 0.11757 | 0.07 | 0.09 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.08 | 0.09 | 0.06 | 0.08 | < 5 | 0.1 | 0.09 | < 0.1 | 0.08875 | 0.1 | 0.18 | 0.07 | 0.1 | 0.09 | 0.11 | [7 empty] | 0.089437
| 1 | 5 | 0.13 | 0.09 | 0.11384 | 0.09 | 0.11 | < 0.05 | < 0.1 | < 5 | < 0.1 | 0.07 | 0.09 | 0.1 | 0.09 | < 5 | 0.09 | 0.08 | < 0.1 | 0.12126 | 0.09 | 0.154 | 0.04 | 0.11 | 0.08 | 0.1 | [7 empty] | 122
| 1 | 6 | 0.06 | 0.09 | 0.12987 | 0.09 | 0.08 | < 0.05 | 0.1 | < 5 | < 0.1 | 0.07 | 0.09 | 0.09 | 0.08 | < 5 | 0.11 | 0.08 | < 0.1 | 0.08047 | 0.09 | 0.257 | < 0.01 | 0.08 | 0.07 | 0.18
| Mean |  | 0.086 | 0.09 | 0.11898 | 0.083333 | 0.095 |  | 0.1 |  | 0.1 | 0.083333 | 0.083333 | 0.083333 | 0.088333 |  | 0.096667 | 0.085 |  | 0.10803 | 0.096667 | 0.317 | 0.044 | 0.095 | 0.083333 | 0.10833
| Median |  | 0.08 | 0.09 | 0.11917 | 0.085 | 0.09 |  | 0.1 |  | 0.1 | 0.085 | 0.085 | 0.09 | 0.085 |  | 0.1 | 0.085 |  | 0.10501 | 0.095 | 0.2955 | 0.04 | 0.095 | 0.085 | 0.105
| Std Dev. |  | 0.029665 | 1.5202e-17 | 0.0075362 | 0.012111 | 0.016432 |  | 0 |  |  | 0.012111 | 0.008165 | 0.015055 | 0.01472 |  | 0.010328 | 0.0054772 |  | 0.034358 | 0.008165 | 0.15057 | 0.015166 | 0.010488 | 0.008165 | 0.040702
| Rel.Std.Dev. |  | 34.49% | 1.689e-14% | 6.334% | 14.53% | 17.3% |  | 0% |  |  | 14.53% | 9.798% | 18.07% | 16.66% |  | 10.68% | 6.444% |  | 31.8% | 8.447% | 47.5% | 34.47% | 11.04% | 9.798% | 37.57%
| PDM3 |  | -3.843% | 0.6295% | 33.03% | -6.825% | 6.22% |  | 11.81% |  | 11.81% | -6.825% | -6.825% | -6.825% | -1.234% |  | 8.084% | -4.961% |  | 20.79% | 8.084% | 254.4% | -50.8% | 6.22% | -6.825% | 21.13%
| Z-Score (Absolute) |  | 1.2 | 0.07 | 2.68 | 0.52 | 0.52 | 5.77 | NA | 216.75 | NA | 0.52 | 0.52 | 0.52 | 0.07 | 216.75 | 0.67 | 0.37 | 3.52 | 1.7 | 0.67 | 20.48 | 4.65 | 0.52 | 0.52 | 1.72
Results from laboratories 1.11 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 233b (Certified Value 0.72 ppm) | [32 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.4 | 0.751 | 0.63 | 0.54549 | 0.81729 | 0.8 | 0.8 | 0.89 | 0.74 | 0.8 | 0.7 | 0.7 | 0.7 | < 100 | 0.7 | 0.67 | 0.6 | 0.8 | 0.7 | 0.7 | 1.7 | 0.7 | 0.7 | 0.7 | [7 empty] | 1
| 1 | 2 | 0.4 | 0.757 | 0.67 | 0.57012 | 0.79498 | 0.8 | 0.78 | 0.8 | 0.73 | 0.8 | 0.7 | 0.7 | 0.7 | < 100 | 0.7 | 0.7 | 0.6 | 0.7 | 0.7 | 0.7 | 1.6 | 0.7 | 0.8 | 0.7 | [7 empty] | 24
| 1 | 3 | 0.4 | 0.749 | 0.621 | 0.57317 | 0.79682 | 0.8 | 0.78 | 0.81 | 0.73 | 0.8 | 0.7 | 0.7 | 0.7 | < 100 | 0.7 | 0.71 | 0.6 | 0.8 | 0.7 | 0.7 | 1.5 | 0.6 | 0.8 | 0.7 | [7 empty] | 16
| 1 | 4 | 0.4 | 0.754 | 0.623 | 0.54636 | 0.79951 | 0.8 | 0.8 | 0.88 | 0.72 | 0.8 | 0.7 | 0.7 | 0.7 | < 100 | 0.7 | 0.67 | 0.6 | 0.8 | 0.7 | 0.7 | 1.5 | 0.7 | 0.7 | 0.7 | [7 empty] | 0.7188
| 1 | 5 | 0.4 | 0.749 | 0.63 | 0.54648 | 0.75497 | 0.8 | 0.77 | 0.81 | 0.73 | 0.8 | 0.7 | 0.7 | 0.7 | < 100 | 0.7 | 0.65 | 0.6 | 0.8 | 0.7 | 0.7 | 1.4 | 0.7 | 0.7 | 0.7 | [7 empty] | 123
| 1 | 6 | 0.4 | 0.769 | 0.637 | 0.56352 | 0.79632 | 0.8 | 0.8 | 0.81 | 0.74 | 0.8 | 0.7 | 0.7 | 0.7 | < 100 | 0.7 | 0.7 | 0.6 | 0.8 | 0.7 | 0.7 | 1.3 | 0.7 | 0.8 | 0.7
| Mean |  | 0.4 | 0.75483 | 0.63517 | 0.55752 | 0.79331 | 0.8 | 0.78833 | 0.83333 | 0.73167 | 0.8 | 0.7 | 0.7 | 0.7 |  | 0.7 | 0.68333 | 0.6 | 0.78333 | 0.7 | 0.7 | 1.5 | 0.68333 | 0.75 | 0.7
| Median |  | 0.4 | 0.7525 | 0.63 | 0.555 | 0.79657 | 0.8 | 0.79 | 0.81 | 0.73 | 0.8 | 0.7 | 0.7 | 0.7 |  | 0.7 | 0.685 | 0.6 | 0.8 | 0.7 | 0.7 | 1.5 | 0.7 | 0.75 | 0.7
| Std Dev. |  | 6.0809e-17 | 0.0076004 | 0.017994 | 0.012889 | 0.020532 | 1.2162e-16 | 0.013292 | 0.040332 | 0.0075277 | 1.2162e-16 | 1.2162e-16 | 1.2162e-16 | 1.2162e-16 |  | 1.2162e-16 | 0.023381 | 0 | 0.040825 | 1.2162e-16 | 1.2162e-16 | 0.14142 | 0.040825 | 0.054772 | 1.2162e-16
| Rel.Std.Dev. |  | 1.52e-14% | 1.007% | 2.833% | 2.312% | 2.588% | 1.52e-14% | 1.686% | 4.84% | 1.029% | 1.52e-14% | 1.737e-14% | 1.737e-14% | 1.737e-14% |  | 1.737e-14% | 3.422% | 0% | 5.212% | 1.737e-14% | 1.737e-14% | 9.428% | 5.974% | 7.303% | 1.737e-14%
| PDM3 |  | -44.35% | 5.012% | -11.64% | -22.44% | 10.37% | 11.3% | 9.673% | 15.93% | 1.79% | 11.3% | -2.616% | -2.616% | -2.616% |  | -2.616% | -4.935% | -16.53% | 8.977% | -2.616% | -2.616% | 108.7% | -4.935% | 4.34% | -2.616%
| Z-Score (Absolute) |  | 3.38 | 0.62 | 0.73 | 1.61 | 1.05 | 1.13 | 1 | 1.5 | 0.36 | 1.13 | 0 | 0 | 0 | 555.6 | 0 | 0.19 | 1.13 | 0.94 | 0 | 0 | 9.02 | 0.19 | 0.56 | 0
| Analytical results for Ti in OREAS 233b (Certified Value 0.313 wt.%) | [32 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.3 | 0.2792 | 0.37582 | 0.42956 | 0.33864 | 0.39 | 0.4148 | 0.431 | 0.3857 | 0.31 | 0.27 | 0.311 | 0.322 | 0.294 | < 0.01 | 0.275 | 0.23 | 0.36841 | 0.3967 | 0.242 | 0.3 | 0.22638 | 0.29973 | 0.116 | 0.293 | 0.265 | 0.3 | 0.273 | 0.213 |  |  | 1
| 1 | 2 | 0.31 | 0.2779 | 0.35587 | 0.43637 | 0.33508 | 0.39 | 0.4196 | 0.428 | 0.3816 | 0.3 | 0.278 | 0.318 | 0.334 | 0.303 | < 0.01 | 0.301 | 0.27 | 0.42617 | 0.3423 | 0.271 | 0.3 | 0.18913 | 0.30517 | 0.1165 | 0.291 | 0.302 | 0.31 | 0.295 | 0.215 |  |  | 25
| 1 | 3 | 0.33 | 0.2873 | 0.33213 | 0.43337 | 0.35955 | 0.37 | 0.4256 | 0.442 | 0.383 | 0.29 | 0.273 | 0.322 | 0.331 | 0.301 | < 0.01 | 0.277 | 0.27 | 0.43933 | 0.3354 | 0.242 | 0.29 | 0.25853 | 0.32512 | 0.107 | 0.286 | 0.288 | 0.3 | 0.282 | 0.214 |  |  | 16
| 1 | 4 | 0.34 | 0.2818 | 0.33346 | 0.43267 | 0.32787 | 0.38 | 0.3873 | 0.435 | 0.3728 | 0.31 | 0.272 | 0.323 | 0.33 | 0.304 | < 0.01 | 0.296 | 0.26 | 0.43157 | 0.3281 | 0.218 | 0.3 | 0.23679 | 0.31468 | 0.1105 | 0.289 | 0.268 | 0.3 | 0.284 | 0.203 |  |  | 0.31347
| 1 | 5 | 0.34 | 0.282 | 0.34586 | 0.43719 | 0.3259 | 0.39 | 0.3933 | 0.446 | 0.3656 | 0.31 | 0.245 | 0.317 | 0.334 | 0.306 | < 0.01 | 0.283 | 0.24 | 0.41624 | 0.3395 | 0.245 | 0.28 | 0.21936 | 0.32595 | 0.101 | 0.279 | 0.25 | 0.32 | 0.278 | 0.219 |  |  | 124
| 1 | 6 | 0.32 | 0.2761 | 0.36596 | 0.43667 | 0.33464 | 0.38 | 0.4268 | 0.433 | 0.385 | 0.31 | 0.252 | 0.317 | 0.339 | 0.308 | < 0.01 | 0.298 | 0.23 | 0.42539 | 0.3435 | 0.241 | 0.27 | 0.24813 | 0.31754 | 0.103 | 0.286 | 0.231 | 0.3 | 0.277 | 0.214
| Mean |  | 0.32333 | 0.28072 | 0.35151 | 0.4343 | 0.33695 | 0.38333 | 0.41123 | 0.43583 | 0.37895 | 0.305 | 0.265 | 0.318 | 0.33167 | 0.30267 |  | 0.28833 | 0.25 | 0.41785 | 0.34758 | 0.24317 | 0.29 | 0.22972 | 0.3147 | 0.109 | 0.28733 | 0.26733 | 0.305 | 0.2815 | 0.213
| Median |  | 0.325 | 0.2805 | 0.35086 | 0.43487 | 0.33486 | 0.385 | 0.4172 | 0.434 | 0.3823 | 0.31 | 0.271 | 0.3175 | 0.3325 | 0.3035 |  | 0.2895 | 0.25 | 0.42578 | 0.3409 | 0.242 | 0.295 | 0.23159 | 0.31611 | 0.10875 | 0.2875 | 0.2665 | 0.3 | 0.28 | 0.214
| Std Dev. |  | 0.01633 | 0.0039423 | 0.01762 | 0.0029747 | 0.01206 | 0.008165 | 0.016887 | 0.0068532 | 0.0080204 | 0.0083666 | 0.013236 | 0.0042895 | 0.0056804 | 0.0048854 |  | 0.011378 | 0.018974 | 0.025388 | 0.0247 | 0.016845 | 0.012649 | 0.024431 | 0.010564 | 0.0065038 | 0.0049261 | 0.025516 | 0.0083666 | 0.0076616 | 0.0053292
| Rel.Std.Dev. |  | 5.05% | 1.404% | 5.013% | 0.6849% | 3.579% | 2.13% | 4.106% | 1.572% | 2.116% | 2.743% | 4.995% | 1.349% | 1.713% | 1.614% |  | 3.946% | 7.589% | 6.076% | 7.106% | 6.928% | 4.362% | 10.64% | 3.357% | 5.967% | 1.714% | 9.545% | 2.743% | 2.722% | 2.502%
| PDM3 |  | 3.147% | -10.45% | 12.14% | 38.55% | 7.49% | 22.29% | 31.19% | 39.04% | 20.89% | -2.701% | -15.46% | 1.446% | 5.805% | -3.446% |  | -8.018% | -20.25% | 33.3% | 10.88% | -22.43% | -7.487% | -26.72% | 0.3926% | -65.23% | -8.337% | -14.72% | -2.701% | -10.2% | -32.05%
| Z-Score (Absolute) |  | 0.31 | 0.41 | 0.78 | 2.18 | 0.54 | 1.32 | 1.79 | 2.21 | 1.25 | 0 | 0.67 | 0.22 | 0.45 | 0.04 | 5.06 | 0.28 | 0.93 | 1.9 | 0.72 | 1.04 | 0.25 | 1.27 | 0.16 | 3.3 | 0.3 | 0.63 | 0 | 0.4 | 1.55
| Analytical results for Tl in OREAS 233b (Certified Value 0.11 ppm) | [32 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.1 | 0.11 | 0.11611 | 0.11 | 0.11 | 0.11 | < 10 | 0.12 | 0.1 | 0.11 | 0.1 | 0.11 | < 5 | 0.11 | 0.11 | 0.12 | 1.95 | 0.1 | 0.10818 | 0.1 | 0.112 | 0.03 | 0.1 | 0.11 | 0.1 | [6 empty] | 1
| 1 | 2 | 0.1 | 0.12 | 0.11183 | 0.11 | 0.11 | 0.11 | < 10 | 0.13 | 0.09 | 0.12 | 0.11 | 0.11 | < 5 | 0.1 | 0.11 | 0.11 | 1.72 | 0.1 | 0.1146 | 0.11 | 0.119 | 0.04 | 0.11 | 0.11 | 0.1 | [6 empty] | 26
| 1 | 3 | 0.1 | 0.11 | 0.15678 | 0.11 | 0.11 | 0.11 | < 10 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | < 5 | 0.11 | 0.1 | 0.12 | 7.09 | 0.1 | 0.11151 | 0.11 | 0.112 | 0.03 | 0.11 | 0.11 | 0.1 | [6 empty] | 16
| 1 | 4 | 0.1 | 0.11 | 0.12073 | 0.11 | 0.12 | 0.1 | < 10 | 0.11 | 0.12 | 0.11 | 0.11 | 0.1 | < 5 | 0.1 | 0.11 | 0.11 | 7.38 | 0.1 | 0.11884 | 0.1 | 0.113 | 0.04 | 0.1 | 0.11 | 0.1 | [6 empty] | 0.10918
| 1 | 5 | 0.11 | 0.11 | 0.11176 | 0.11 | 0.12 | 0.11 | < 10 | 0.12 | 0.1 | 0.11 | 0.11 | 0.11 | < 5 | 0.1 | 0.1 | 0.12 | 1.23 | 0.1 | 0.10956 | 0.11 | 0.109 | 0.03 | 0.1 | 0.11 | 0.1 | [6 empty] | 125
| 1 | 6 | 0.11 | 0.11 | 0.12531 | 0.11 | 0.11 | 0.11 | < 10 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | < 5 | 0.11 | 0.1 | 0.11 | 5.11 | 0.1 | 0.11032 | 0.11 | 0.11 | 0.02 | 0.1 | 0.11 | 0.1
| Mean |  | 0.10333 | 0.11167 | 0.12375 | 0.11 | 0.11333 | 0.10833 |  | 0.12 | 0.105 | 0.11167 | 0.10833 | 0.10833 |  | 0.105 | 0.105 | 0.115 | 4.08 | 0.1 | 0.11217 | 0.10667 | 0.1125 | 0.031667 | 0.10333 | 0.11 | 0.1
| Median |  | 0.1 | 0.11 | 0.11842 | 0.11 | 0.11 | 0.11 |  | 0.12 | 0.105 | 0.11 | 0.11 | 0.11 |  | 0.105 | 0.105 | 0.115 | 3.53 | 0.1 | 0.11091 | 0.11 | 0.112 | 0.03 | 0.1 | 0.11 | 0.1
| Std Dev. |  | 0.005164 | 0.0040825 | 0.017009 | 0 | 0.005164 | 0.0040825 |  | 0.0063246 | 0.010488 | 0.0040825 | 0.0040825 | 0.0040825 |  | 0.0054772 | 0.0054772 | 0.0054772 | 2.8014 | 1.5202e-17 | 0.0039239 | 0.005164 | 0.0035071 | 0.0075277 | 0.005164 | 0 | 1.5202e-17
| Rel.Std.Dev. |  | 4.997% | 3.656% | 13.74% | 0% | 4.556% | 3.768% |  | 5.27% | 9.989% | 3.656% | 3.768% | 3.768% |  | 5.216% | 5.216% | 4.763% | 68.66% | 1.52e-14% | 3.498% | 4.841% | 3.117% | 23.77% | 4.997% | 0% | 1.52e-14%
| PDM3 |  | -5.351% | 2.282% | 13.35% | 0.755% | 3.808% | -0.7716% |  | 9.915% | -3.825% | 2.282% | -0.7716% | -0.7716% |  | -3.825% | -3.825% | 5.335% | 3637% | -8.405% | 2.741% | -2.298% | 3.045% | -70.99% | -5.351% | 0.755% | -8.405%
| Z-Score (Absolute) |  | 0.9 | 0.22 | 1.85 | 0 | 0.45 | 0.22 | 659.47 | 1.35 | 0.67 | 0.22 | 0.22 | 0.22 | 322.32 | 0.67 | 0.67 | 0.67 | 535.4 | 1.35 | 0.29 | 0.45 | 0.34 | 10.56 | 0.9 | 0 | NA
Results from laboratory 1.34 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Tm in OREAS 233b (Certified Value 0.16 ppm) | [32 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.05: Lab; 1.07: Lab; 1.10: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.05: 1.04; 1.07: 1.07; 1.10: 1.27 | [6 empty] | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.05: AR*MS; 1.07: AR*OES/MS; 1.10: AR*MS | [6 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.05: 0.5g; 1.07: 0.5g; 1.10: 25g | [6 empty] | 2
| 1 | 1 | 1.01: 0.164; 1.02: 0.164; 1.05: 0.15383; 1.07: 0.2; 1.10: 0.15 | [6 empty] | 1
| 1 | 2 | 1.01: 0.163; 1.02: 0.166; 1.05: 0.1507; 1.07: 0.2; 1.10: 0.15 | [6 empty] | 27
| 1 | 3 | 1.01: 0.167; 1.02: 0.146; 1.05: 0.15314; 1.07: 0.2; 1.10: 0.15 | [6 empty] | 16
| 1 | 4 | 1.01: 0.169; 1.02: 0.149; 1.05: 0.15165; 1.07: 0.2; 1.10: 0.1 | [6 empty] | 0.15996
| 1 | 5 | 1.01: 0.169; 1.02: 0.155; 1.05: 0.1467; 1.07: 0.2; 1.10: 0.1 | [6 empty] | 126
| 1 | 6 | 1.01: 0.165; 1.02: 0.165; 1.05: 0.15078; 1.07: 0.2; 1.10: 0.1
| Mean |  | 1.01: 0.16617; 1.02: 0.1575; 1.05: 0.15113; 1.07: 0.2; 1.10: 0.125
| Median |  | 1.01: 0.166; 1.02: 0.1595; 1.05: 0.15121; 1.07: 0.2; 1.10: 0.125
| Std Dev. |  | 1.01: 0.0025626; 1.02: 0.008735; 1.05: 0.0025091; 1.07: 3.0405e-17; 1.10: 0.027386
| Rel.Std.Dev. |  | 1.01: 1.542%; 1.02: 5.546%; 1.05: 1.66%; 1.07: 1.52e-14%; 1.10: 21.91%
| PDM3 |  | 1.01: 3.88%; 1.02: -1.538%; 1.05: -5.518%; 1.07: 25.03%; 1.10: -21.86%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0; 1.05: 0.5; 1.07: 3.31; 1.10: 2.53
| Analytical results for U in OREAS 233b (Certified Value 0.19 ppm) | [32 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.19 | 0.189 | 0.213 | 0.14853 | 0.22018 | 0.2 | 0.67 | 0.24 | 0.21 | 0.2 | 0.18 | 0.2 | 0.18 | 1552.113 | 0.18 | < 10 | 0.19 | 0.2 | 0.19 | 0.2 | 0.18 | 0.173 | 0.17 | 0.18 | 0.19 | 0.17 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.197 | 0.196 | 0.1608 | 0.21779 | 0.2 | 0.27 | 0.24 | 0.2 | 0.2 | 0.19 | 0.21 | 0.18 | 1587.631 | 0.19 | < 10 | 0.19 | 0.2 | 0.19 | 0.2 | 0.18 | 0.175 | 0.18 | 0.19 | 0.2 | 0.17 | [5 empty] | 28
| 1 | 3 | 0.2 | 0.199 | 0.185 | 0.15727 | 0.25946 | 0.2 | 0.32 | 0.25 | 0.21 | 0.2 | 0.19 | 0.2 | 0.19 | 1673.437 | 0.18 | < 10 | 0.2 | 0.2 | 0.19 | 0.2 | 0.18 | 0.181 | 0.18 | 0.19 | 0.19 | 0.17 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.196 | 0.185 | 0.14505 | 0.2327 | 0.2 | 0.37 | 0.26 | 0.2 | 0.2 | 0.17 | 0.19 | 0.19 | 1695.68 | 0.19 | < 10 | 0.18 | 0.2 | 0.19 | 0.2 | 0.18 | 0.171 | 0.17 | 0.19 | 0.19 | 0.17 | [5 empty] | 0.1896
| 1 | 5 | 0.2 | 0.194 | 0.189 | 0.15425 | 0.21854 | 0.2 | 0.28 | 0.26 | 0.21 | 0.2 | 0.19 | 0.19 | 0.18 | 1426.684 | 0.19 | < 10 | 0.18 | 0.2 | 0.2 | 0.2 | 0.18 | 0.172 | 0.17 | 0.18 | 0.19 | 0.17 | [5 empty] | 127
| 1 | 6 | 0.2 | 0.201 | 0.201 | 0.15732 | 0.22025 | 0.2 | 0.39 | 0.25 | 0.21 | 0.2 | 0.18 | 0.2 | 0.19 | 1467.193 | 0.19 | < 10 | 0.2 | 0.2 | 0.19 | 0.2 | 0.18 | 0.177 | 0.16 | 0.18 | 0.19 | 0.17
| Mean |  | 0.19833 | 0.196 | 0.19483 | 0.15387 | 0.22815 | 0.2 | 0.38333 | 0.25 | 0.20667 | 0.2 | 0.18333 | 0.19833 | 0.185 | 1567.123 | 0.18667 |  | 0.19 | 0.2 | 0.19167 | 0.2 | 0.18 | 0.17483 | 0.17167 | 0.185 | 0.19167 | 0.17
| Median |  | 0.2 | 0.1965 | 0.1925 | 0.15576 | 0.22021 | 0.2 | 0.345 | 0.25 | 0.21 | 0.2 | 0.185 | 0.2 | 0.185 | 1569.872 | 0.19 |  | 0.19 | 0.2 | 0.19 | 0.2 | 0.18 | 0.174 | 0.17 | 0.185 | 0.19 | 0.17
| Std Dev. |  | 0.0040825 | 0.0041952 | 0.010926 | 0.0059662 | 0.01629 | 3.0405e-17 | 0.14828 | 0.0089443 | 0.005164 | 3.0405e-17 | 0.008165 | 0.0075277 | 0.0054772 | 107.8745 | 0.005164 |  | 0.0089443 | 3.0405e-17 | 0.0040825 | 3.0405e-17 | 3.0405e-17 | 0.0037103 | 0.0075277 | 0.0054772 | 0.0040825 | 0
| Rel.Std.Dev. |  | 2.058% | 2.14% | 5.608% | 3.877% | 7.14% | 1.52e-14% | 38.68% | 3.578% | 2.499% | 1.52e-14% | 4.454% | 3.795% | 2.961% | 6.884% | 2.766% |  | 4.708% | 1.52e-14% | 2.13% | 1.52e-14% | 1.689e-14% | 2.122% | 4.385% | 2.961% | 2.13% | 0%
| PDM3 |  | 4.607% | 3.376% | 2.761% | -18.84% | 20.33% | 5.486% | 102.2% | 31.86% | 9.002% | 5.486% | -3.305% | 4.607% | -2.426% | 826444% | -1.547% |  | 0.2113% | 5.486% | 1.09% | 5.486% | -5.063% | -7.788% | -9.458% | -2.426% | 1.09% | -10.34%
| Z-Score (Absolute) |  | 0.45 | 0.29 | 0.21 | 2.55 | 2.46 | NA | 12.92 | 3.93 | 1.01 | NA | 0.56 | 0.45 | 0.45 | 105659.56 | 0.34 | 324.23 | 0.11 | NA | 0 | NA | 0.79 | 1.14 | 1.35 | 0.45 | 0 | 1.46
Results from laboratories 1.07, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 233b (Certified Value 131 ppm) | [32 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 128 | 133.1 | 128 | 145.8043 | 144.9919 | 149 | 164 | 176 | 164 | 133 | 138 | 128 | 131 | 122 | < 10 | 122 | 110 | 154.6141 | 149 | 113 | 133 | 105.98 | 130.4744 | 90 | 125 | 112 | 122 | 128 | 98.8 |  |  | 1
| 1 | 2 | 135 | 134.4 | 122.6 | 149.8542 | 142.8498 | 145 | 164 | 173 | 162 | 131 | 138 | 130 | 135 | 124 | < 10 | 127 | 123 | 171.7385 | 130 | 125 | 133 | 94.15 | 132.7376 | 90 | 124 | 123 | 121 | 135 | 98.7 |  |  | 29
| 1 | 3 | 138 | 135.4 | 117.05 | 149.5593 | 153.9666 | 143 | 159 | 177 | 165 | 131 | 133 | 130 | 133 | 124 | < 10 | 121 | 121 | 176.75 | 134 | 118 | 135 | 114.67 | 141.2752 | 85 | 123 | 119 | 121 | 131 | 96.8 |  |  | 16
| 1 | 4 | 140 | 136.2 | 116.9 | 147.9321 | 140.9228 | 139 | 166 | 175 | 162 | 132 | 134 | 131 | 132 | 125 | < 10 | 126 | 117 | 174.3146 | 136 | 112 | 127 | 109.03 | 144.7896 | 85 | 124 | 117 | 122 | 131 | 93.9 |  |  | 130.6089
| 1 | 5 | 139 | 134.9 | 120.96 | 144.898 | 140.1337 | 145 | 160 | 180 | 161 | 133 | 128 | 130 | 133 | 126 | < 10 | 125 | 111 | 170.4878 | 140 | 114 | 128 | 103.07 | 142.37 | 85 | 121 | 109 | 121 | 130 | 97.7 |  |  | 128
| 1 | 6 | 135 | 133 | 126.62 | 148.2934 | 143.4365 | 148 | 169 | 175 | 165 | 135 | 129 | 130 | 135 | 127 | < 10 | 129 | 108 | 173.2755 | 147 | 120 | 120 | 112.4 | 139.4168 | 85 | 123 | 105 | 123 | 129 | 95.8
| Mean |  | 135.8333 | 134.5 | 122.0217 | 147.7235 | 144.3835 | 144.8333 | 163.6667 | 176 | 163.1667 | 132.5 | 133.3333 | 129.8333 | 133.1667 | 124.6667 |  | 125 | 115 | 170.1968 | 139.3333 | 117 | 129.3333 | 106.55 | 138.5106 | 86.6667 | 123.3333 | 114.1667 | 121.6667 | 130.6667 | 96.95
| Median |  | 136.5 | 134.65 | 121.78 | 148.1128 | 143.1431 | 145 | 164 | 175.5 | 163 | 132.5 | 133.5 | 130 | 133 | 124.5 |  | 125.5 | 114 | 172.507 | 138 | 116 | 130.5 | 107.505 | 140.346 | 85 | 123.5 | 114.5 | 121.5 | 130.5 | 97.25
| Std Dev. |  | 4.3551 | 1.2712 | 4.6754 | 1.9973 | 5.0098 | 3.6009 | 3.7238 | 2.3664 | 1.7224 | 1.5166 | 4.274 | 0.98319 | 1.6021 | 1.7512 |  | 3.0332 | 6.229 | 7.933 | 7.4744 | 4.98 | 5.5377 | 7.3824 | 5.6691 | 2.582 | 1.3663 | 6.7057 | 0.8165 | 2.4221 | 1.8812
| Rel.Std.Dev. |  | 3.206% | 0.9451% | 3.832% | 1.352% | 3.47% | 2.486% | 2.275% | 1.345% | 1.056% | 1.145% | 3.205% | 0.7573% | 1.203% | 1.405% |  | 2.427% | 5.416% | 4.661% | 5.364% | 4.256% | 4.282% | 6.929% | 4.093% | 2.979% | 1.108% | 5.874% | 0.6711% | 1.854% | 1.94%
| PDM3 |  | 4% | 2.979% | -6.575% | 13.1% | 10.55% | 10.89% | 25.31% | 34.75% | 24.93% | 1.448% | 2.086% | -0.5938% | 1.958% | -4.55% |  | -4.294% | -11.95% | 30.31% | 6.68% | -10.42% | -0.9766% | -18.42% | 6.05% | -33.64% | -5.571% | -12.59% | -6.847% | 0.04422% | -25.77%
| Z-Score (Absolute) |  | 0.39 | 0.29 | 0.65 | 1.28 | 1.03 | 1.06 | 2.47 | 3.4 | 2.44 | 0.14 | 0.2 | 0.06 | 0.19 | 0.45 | 9.42 | 0.42 | 1.17 | 2.96 | 0.65 | 1.02 | 0.1 | 1.81 | 0.59 | 3.3 | 0.55 | 1.24 | 0.67 | 0 | 2.53
| Analytical results for W in OREAS 233b (Certified Value 21.4 ppm) | [32 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 23.8 | 19.24 | 27.1 | 26.8168 | 18.7 | 24 | 15.7 | 16.7 | 18 | 22.2 | 22.4 | 22.4 | 21.4 | < 100 | 21.6 | 16 | 21.2 | 15.2 | 24 | 18 | 23.3 | 25.1 | 23.6 | 18.76 | 26 | [6 empty] | 1
| 1 | 2 | 23.8 | 19.08 | 26.33 | 27.3058 | 18.4 | 23 | 15.2 | 16.88 | 18 | 24.6 | 21.7 | 23.4 | 22.3 | < 100 | 22.5 | 17 | 22.5 | 15.7 | 23 | 18 | 22.8 | 26.2 | 24.5 | 19.51 | 26.2 | [6 empty] | 30
| 1 | 3 | 24 | 19.22 | 26.12 | 31.2242 | 15 | 23 | 15.2 | 15.96 | 18 | 22 | 23.1 | 23.1 | 22.1 | < 100 | 21.8 | 17 | 21.9 | 15.3 | 23 | 18 | 23 | 26.2 | 23.9 | 18.85 | 25.9 | [6 empty] | 16
| 1 | 4 | 23.8 | 19.04 | 28.03 | 27.6556 | 17.3 | 25 | 17.3 | 15.44 | 18 | 22.5 | 23 | 22 | 23.2 | < 100 | 21.4 | 17 | 21.8 | 16.4 | 24 | 18 | 22.9 | 26.4 | 24 | 19 | 26.1 | [6 empty] | 21.4458
| 1 | 5 | 24.4 | 19.1 | 26.5 | 25.4293 | 18.6 | 26 | 15.1 | 16.73 | 18 | 22.9 | 22.3 | 23.2 | 22.7 | < 100 | 21.6 | 16 | 21.2 | 15.1 | 24 | 19 | 23 | 26 | 24 | 19.06 | 25.9 | [6 empty] | 129
| 1 | 6 | 24.6 | 19.5 | 27.3 | 26.8771 | 19.1 | 25 | 14.9 | 16.38 | 18 | 23.1 | 22.5 | 22.7 | 22.4 | < 100 | 21.7 | 16 | 22.1 | 15.5 | 23 | 19 | 22.3 | 25.4 | 23.1 | 19.21 | 25.4
| Mean |  | 24.0667 | 19.1967 | 26.8967 | 27.5515 | 17.85 | 24.3333 | 15.5667 | 16.3483 | 18 | 22.8833 | 22.5 | 22.8 | 22.35 |  | 21.7667 | 16.5 | 21.7833 | 15.5333 | 23.5 | 18.3333 | 22.8833 | 25.8833 | 23.85 | 19.065 | 25.9167
| Median |  | 23.9 | 19.16 | 26.8 | 27.0914 | 18.5 | 24.5 | 15.2 | 16.54 | 18 | 22.7 | 22.45 | 22.9 | 22.35 |  | 21.65 | 16.5 | 21.85 | 15.4 | 23.5 | 18 | 22.95 | 26.1 | 23.95 | 19.03 | 25.95
| Std Dev. |  | 0.35024 | 0.16848 | 0.71696 | 1.9524 | 1.5215 | 1.2111 | 0.88919 | 0.55297 | 0 | 0.93684 | 0.5099 | 0.53292 | 0.60249 |  | 0.38297 | 0.54772 | 0.51153 | 0.4761 | 0.54772 | 0.5164 | 0.33116 | 0.51543 | 0.46797 | 0.26928 | 0.27869
| Rel.Std.Dev. |  | 1.455% | 0.8777% | 2.666% | 7.086% | 8.524% | 4.977% | 5.712% | 3.382% | 0% | 4.094% | 2.266% | 2.337% | 2.696% |  | 1.759% | 3.32% | 2.348% | 3.065% | 2.331% | 2.817% | 1.447% | 1.991% | 1.962% | 1.412% | 1.075%
| PDM3 |  | 12.22% | -10.49% | 25.42% | 28.47% | -16.77% | 13.46% | -27.41% | -23.77% | -16.07% | 6.703% | 4.915% | 6.314% | 4.216% |  | 1.496% | -23.06% | 1.574% | -27.57% | 9.578% | -14.51% | 6.703% | 20.69% | 11.21% | -11.1% | 20.85%
| Z-Score (Absolute) |  | 0.31 | 0.66 | 0.88 | 1.01 | 0.93 | 0.37 | 1.38 | 1.23 | 0.9 | 0.08 | 0 | 0.06 | 0.03 | 5.48 | 0.15 | 1.2 | 0.14 | 1.39 | 0.2 | 0.83 | 0.08 | 0.67 | 0.27 | 0.68 | 0.68
| Analytical results for Y in OREAS 233b (Certified Value 10.7 ppm) | [32 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 10.17 | 11.11 | 10.73 | 10.5882 | 10.9454 | 11.2 | 10.8 | 13.6 | 10.75 | 9 | 10.9 | 10.3 | 10.65 | 9.88 | 10.45 | 9 | 11.37 | 10.7 | 8.8 | 11.2 | 9.44 | 9.8 | 11.01 | 9.38 | [7 empty] | 1
| 1 | 2 | 10.53 | 11.07 | 10.48 | 10.5052 | 10.9626 | 11.3 | 11 | 13.4 | 10.76 | 9 | 11 | 10.35 | 10.5 | 10.15 | 10.6 | 10 | 10.92 | 10.5 | 8.7 | 11.25 | 10 | 10.1 | 11.34 | 9.42 | [7 empty] | 31
| 1 | 3 | 10.75 | 11.22 | 10.23 | 10.6176 | 11.8377 | 11.2 | 10.9 | 13.8 | 10.95 | 9 | 10.7 | 10.7 | 10.3 | 10.25 | 10.3 | 10 | 11.26 | 10.5 | 8.4 | 11.25 | 9.69 | 10.4 | 11.22 | 9.26 | [7 empty] | 16
| 1 | 4 | 10.65 | 11.22 | 10.17 | 10.6339 | 10.8299 | 11 | 11.2 | 13.8 | 10.63 | 9 | 10.8 | 10.3 | 10.7 | 10.35 | 10.55 | 9 | 10.63 | 10.4 | 8.4 | 11.2 | 9.62 | 10.3 | 11.13 | 9.06 | [7 empty] | 10.6648
| 1 | 5 | 10.57 | 11.12 | 10.29 | 10.5432 | 10.7651 | 11.2 | 10.4 | 13.9 | 10.44 | 10 | 10.6 | 10.65 | 10.7 | 10.25 | 10.85 | 9 | 11.1 | 10.2 | 8.1 | 11.4 | 9.52 | 10.2 | 10.99 | 9.28 | [7 empty] | 130
| 1 | 6 | 10.51 | 11.06 | 10.65 | 10.5781 | 10.993 | 11.5 | 11.1 | 13.5 | 10.89 | 10 | 10.8 | 10.45 | 10.65 | 10.85 | 10.95 | 9 | 10.51 | 10.2 | 8.2 | 11.15 | 9.11 | 9.9 | 10.98 | 9.09
| Mean |  | 10.53 | 11.1333 | 10.425 | 10.5777 | 11.0556 | 11.2333 | 10.9 | 13.6667 | 10.7367 | 9.3333 | 10.8 | 10.4583 | 10.5833 | 10.2883 | 10.6167 | 9.3333 | 10.965 | 10.4167 | 8.4333 | 11.2417 | 9.5633 | 10.1167 | 11.1117 | 9.2483
| Median |  | 10.55 | 11.115 | 10.385 | 10.5831 | 10.954 | 11.2 | 10.95 | 13.7 | 10.755 | 9 | 10.8 | 10.4 | 10.65 | 10.25 | 10.575 | 9 | 11.01 | 10.45 | 8.4 | 11.225 | 9.57 | 10.15 | 11.07 | 9.27
| Std Dev. |  | 0.19718 | 0.070899 | 0.2315 | 0.047557 | 0.39289 | 0.1633 | 0.28284 | 0.19664 | 0.18392 | 0.5164 | 0.14142 | 0.17725 | 0.15706 | 0.31877 | 0.24427 | 0.5164 | 0.34367 | 0.19408 | 0.27325 | 0.08612 | 0.29412 | 0.23166 | 0.14607 | 0.1473
| Rel.Std.Dev. |  | 1.873% | 0.6368% | 2.221% | 0.4496% | 3.554% | 1.454% | 2.595% | 1.439% | 1.713% | 5.533% | 1.309% | 1.695% | 1.484% | 3.098% | 2.301% | 5.533% | 3.134% | 1.863% | 3.24% | 0.7661% | 3.075% | 2.29% | 1.315% | 1.593%
| PDM3 |  | -1.264% | 4.393% | -2.249% | -0.8173% | 3.664% | 5.331% | 2.205% | 28.15% | 0.6736% | -12.48% | 1.267% | -1.936% | -0.7642% | -3.53% | -0.4516% | -12.48% | 2.815% | -2.327% | -20.92% | 5.409% | -10.33% | -5.14% | 4.19% | -13.28%
| Z-Score (Absolute) |  | 0.14 | 1.06 | 0.35 | 0.04 | 0.91 | 1.26 | 0.6 | 6.11 | 0.27 | NA | 0.4 | 0.28 | 0.03 | 0.62 | 0.03 | NA | 0.73 | 0.37 | 4.32 | 1.28 | 2.07 | 0.96 | 1.02 | 2.69
Results from laboratories 1.12 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 233b (Certified Value 0.98 ppm) | [32 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.05: Lab; 1.07: Lab; 1.08: Lab; 1.10: Lab; 1.11: Lab | [6 empty] | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.05: 1.04; 1.07: 1.07; 1.08: 1.08; 1.10: 1.27; 1.11: 1.32 | [6 empty] | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.05: AR*MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.10: AR*MS; 1.11: AR*OES/MS | [6 empty] | 2
|  |  | 1.01: 0.25g; 1.02: 0.5g; 1.03: 0.5g; 1.05: 0.5g; 1.07: 0.5g; 1.08: 0.5g; 1.10: 25g; 1.11: 0.25g | [6 empty] | 2
| 1 | 1 | 1.01: 0.9; 1.02: 1.062; 1.03: 0.887; 1.05: 0.97207; 1.07: 1.2; 1.08: 1.23; 1.10: 0.8; 1.11: 0.9 | [6 empty] | 1
| 1 | 2 | 1.01: 1; 1.02: 1.026; 1.03: 0.88; 1.05: 0.97033; 1.07: 1.1; 1.08: 1.25; 1.10: 0.8; 1.11: 1 | [6 empty] | 32
| 1 | 3 | 1.01: 1; 1.02: 1.094; 1.03: 0.815; 1.05: 0.94755; 1.07: 1.1; 1.08: 1.2; 1.10: 0.75; 1.11: 0.9 | [6 empty] | 16
| 1 | 4 | 1.01: 1; 1.02: 1.051; 1.03: 0.825; 1.05: 0.91937; 1.07: 1.1; 1.08: 1.24; 1.10: 0.75; 1.11: 0.8 | [6 empty] | 0.98243
| 1 | 5 | 1.01: 1; 1.02: 1.085; 1.03: 0.836; 1.05: 0.93203; 1.07: 1.1; 1.08: 1.19; 1.10: 0.7; 1.11: 0.9 | [6 empty] | 131
| 1 | 6 | 1.01: 1; 1.02: 1.054; 1.03: 0.914; 1.05: 0.95606; 1.07: 1.2; 1.08: 1.22; 1.10: 0.7; 1.11: 0.9
| Mean |  | 1.01: 0.98333; 1.02: 1.062; 1.03: 0.8595; 1.05: 0.94957; 1.07: 1.1333; 1.08: 1.2217; 1.10: 0.75; 1.11: 0.9
| Median |  | 1.01: 1; 1.02: 1.058; 1.03: 0.858; 1.05: 0.9518; 1.07: 1.1; 1.08: 1.225; 1.10: 0.75; 1.11: 0.9
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.024633; 1.03: 0.039672; 1.05: 0.020991; 1.07: 0.05164; 1.08: 0.023166; 1.10: 0.044721; 1.11: 0.063246
| Rel.Std.Dev. |  | 1.01: 4.152%; 1.02: 2.32%; 1.03: 4.616%; 1.05: 2.211%; 1.07: 4.556%; 1.08: 1.896%; 1.10: 5.963%; 1.11: 7.027%
| PDM3 |  | 1.01: 0.09245%; 1.02: 8.1%; 1.03: -12.51%; 1.05: -3.345%; 1.07: 15.36%; 1.08: 24.35%; 1.10: -23.66%; 1.11: -8.39%
| Z-Score (Absolute) |  | 1.01: 0.11; 1.02: 0.64; 1.03: 0.71; 1.05: 0.11; 1.07: 1.11; 1.08: 1.7; 1.10: 1.44; 1.11: 0.44
| Analytical results for Zn in OREAS 233b (Certified Value 123 ppm) | [32 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 125 | 122.7 | 111.93 | 118.94 | 128.2453 | 125 | 126 | 129 | 126 | 115 | 136 | 128 | 121 | 126 | < 100 | 137 | 112 | 121.8391 | 122 | 119 | 121 | 117.48 | 120.34 | 120 | 136.9621 | 128 | 106.832 | 119 | 129 | 117 | 126 | 1
| 1 | 2 | 126 | 121.9 | 109.3 | 118.55 | 127.8224 | 126 | 125 | 126 | 126 | 117 | 151 | 128 | 120 | 128 | < 100 | 137 | 118 | 122.0756 | 128 | 127.2 | 121 | 116.56 | 123.72 | 115 | 138.2335 | 127 | 112.175 | 121 | 134 | 118 | 126 | 33
| 1 | 3 | 126 | 122.5 | 109.68 | 117.56 | 131.7541 | 123 | 123 | 128 | 126 | 115 | 132 | 127 | 122 | 128 | < 100 | 135 | 115 | 124.4552 | 136 | 124.3 | 119 | 116.17 | 123.71 | 115 | 133.8584 | 128 | 112.825 | 119 | 133 | 118 | 126 | 16
| 1 | 4 | 127 | 123.6 | 108.19 | 118.77 | 127.9832 | 117 | 124 | 124 | 121 | 117 | 134 | 128 | 121 | 129 | < 100 | 137 | 113 | 123.7789 | 126 | 121.9 | 119 | 118.95 | 123.76 | 115 | 140.2706 | 126 | 105.893 | 124 | 132 | 116 | 124 | 123.4355
| 1 | 5 | 126 | 121.9 | 108.67 | 121.13 | 126.8568 | 123 | 122 | 129 | 124 | 118 | 136 | 128 | 124 | 128 | < 100 | 142 | 113 | 120.6175 | 132 | 119.7 | 120 | 118.69 | 122.91 | 115 | 141.0302 | 127 | 109.778 | 119 | 132 | 118 | 126 | 132
| 1 | 6 | 126 | 120.6 | 109.06 | 115.32 | 129.2771 | 125 | 123 | 127 | 125 | 119 | 135 | 126 | 123 | 129 | < 100 | 141 | 115 | 123.3026 | 121 | 122.3 | 119 | 118.19 | 123.72 | 115 | 132.8693 | 127 | 108.348 | 122 | 128 | 117 | 123
| Mean |  | 126 | 122.2 | 109.4717 | 118.3783 | 128.6565 | 123.1667 | 123.8333 | 127.1667 | 124.6667 | 116.8333 | 137.3333 | 127.5 | 121.8333 | 128 |  | 138.1667 | 114.3333 | 122.6781 | 127.5 | 122.4 | 119.8333 | 117.6733 | 123.0267 | 115.8333 | 137.204 | 127.1667 | 109.3085 | 120.6667 | 131.3333 | 117.3333 | 125.1667
| Median |  | 126 | 122.2 | 109.18 | 118.66 | 128.1143 | 124 | 123.5 | 127.5 | 125.5 | 117 | 135.5 | 128 | 121.5 | 128 |  | 137 | 114 | 122.6891 | 127 | 122.1 | 119.5 | 117.835 | 123.715 | 115 | 137.5978 | 127 | 109.063 | 120 | 132 | 117.5 | 126
| Std Dev. |  | 0.63246 | 1.004 | 1.3093 | 1.9023 | 1.7051 | 3.2506 | 1.472 | 1.9408 | 1.9664 | 1.6021 | 6.8605 | 0.83666 | 1.472 | 1.0954 |  | 2.7142 | 2.1602 | 1.4193 | 5.7879 | 3.0252 | 0.98319 | 1.1373 | 1.3563 | 2.0412 | 3.321 | 0.75277 | 2.8116 | 2.0656 | 2.3381 | 0.8165 | 1.3292
| Rel.Std.Dev. |  | 0.5019% | 0.8216% | 1.196% | 1.607% | 1.325% | 2.639% | 1.189% | 1.526% | 1.577% | 1.371% | 4.996% | 0.6562% | 1.208% | 0.8558% |  | 1.964% | 1.889% | 1.157% | 4.54% | 2.472% | 0.8205% | 0.9665% | 1.102% | 1.762% | 2.42% | 0.592% | 2.572% | 1.712% | 1.78% | 0.6959% | 1.062%
| PDM3 |  | 2.078% | -1.001% | -11.31% | -4.097% | 4.23% | -0.2178% | 0.3223% | 3.023% | 0.9974% | -5.349% | 11.26% | 3.293% | -1.298% | 3.698% |  | 11.93% | -7.374% | -0.6135% | 3.293% | -0.8389% | -2.918% | -4.668% | -0.3312% | -6.159% | 11.15% | 3.023% | -11.44% | -2.243% | 6.398% | -4.944% | 1.403%
| Z-Score (Absolute) |  | 0.45 | 0.12 | 2.04 | 0.7 | 0.85 | 0.02 | 0.12 | 0.62 | 0.25 | 0.93 | 2.16 | 0.67 | 0.18 | 0.75 | 11.01 | 2.28 | 1.31 | 0.05 | 0.67 | 0.09 | 0.48 | 0.81 | 0 | 1.08 | 2.14 | 0.62 | 2.07 | 0.36 | 1.25 | 0.86 | 0.32
| Analytical results for Zr in OREAS 233b (Certified Value 13.8 ppm) | [32 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 11.7 | 14.02 | 20.11 | 15.0609 | 11.9 | 13 | 16.7 | 18 | 12 | 11.8 | 13.6 | 14.8 | 12.5 | 13.2 | 12 | 13 | 12.9 | 15.831 | 7 | 15 | 11 | 12.5 | 14.2 | 11 | [7 empty] | 1
| 1 | 2 | 11.7 | 13.96 | 19.41 | 14.8151 | 12.6 | 12.9 | 16.4 | 18 | 12 | 11.8 | 13.4 | 15.6 | 13.4 | 13.9 | 13 | 13.5 | 12.7 | 16.362 | 8 | 14.9 | 12.8 | 13 | 14.8 | 10.9 | [7 empty] | 34
| 1 | 3 | 12 | 14.34 | 18.66 | 15.9303 | 9.6 | 12.8 | 16.9 | 18.1 | 12 | 11.1 | 14.1 | 15.6 | 13.2 | 13.7 | 13 | 13.3 | 12.6 | 17.244 | 7 | 14.6 | 12.6 | 13.5 | 14.1 | 10.7 | [7 empty] | 16
| 1 | 4 | 13.2 | 14.51 | 18.69 | 14.8105 | 12.1 | 12.7 | 16.3 | 17.4 | 12 | 11.6 | 13.7 | 14.2 | 13.7 | 13.9 | 13 | 12.6 | 12.3 | 17.802 | 8 | 15.2 | 12.1 | 13.1 | 14.6 | 10.4 | [7 empty] | 13.7535
| 1 | 5 | 13.3 | 14.19 | 19.18 | 14.3259 | 12.1 | 12.2 | 15.9 | 17.4 | 12 | 11 | 13.8 | 15.1 | 13.4 | 13.5 | 12 | 12.4 | 11.4 | 17.316 | 6 | 14.5 | 11.9 | 13.4 | 14.1 | 10.9 | [7 empty] | 133
| 1 | 6 | 12.5 | 14.12 | 20.11 | 14.7921 | 13.1 | 12.4 | 16.4 | 18.2 | 12 | 11.6 | 13.7 | 16.1 | 13.8 | 13.8 | 12 | 13.5 | 11.3 | 17.001 | 7 | 14.5 | 10.9 | 12.7 | 14.2 | 10.7
| Mean |  | 12.4 | 14.19 | 19.36 | 14.9558 | 11.9 | 12.6667 | 16.4333 | 17.85 | 12 | 11.4833 | 13.7167 | 15.2333 | 13.3333 | 13.6667 | 12.5 | 13.05 | 12.2 | 16.926 | 7.1667 | 14.7833 | 11.8833 | 13.0333 | 14.3333 | 10.7667
| Median |  | 12.25 | 14.155 | 19.295 | 14.8128 | 12.1 | 12.75 | 16.4 | 18 | 12 | 11.6 | 13.7 | 15.35 | 13.4 | 13.75 | 12.5 | 13.15 | 12.45 | 17.1225 | 7 | 14.75 | 12 | 13.05 | 14.2 | 10.8
| Std Dev. |  | 0.72111 | 0.20572 | 0.64789 | 0.53389 | 1.2083 | 0.30768 | 0.34448 | 0.35637 | 0 | 0.34881 | 0.23166 | 0.67725 | 0.46332 | 0.27325 | 0.54772 | 0.46797 | 0.68702 | 0.71321 | 0.75277 | 0.29269 | 0.79352 | 0.38816 | 0.29439 | 0.21602
| Rel.Std.Dev. |  | 5.815% | 1.45% | 3.347% | 3.57% | 10.15% | 2.429% | 2.096% | 1.996% | 0% | 3.038% | 1.689% | 4.446% | 3.475% | 1.999% | 4.382% | 3.586% | 5.631% | 4.214% | 10.5% | 1.98% | 6.678% | 2.978% | 2.054% | 2.006%
| PDM3 |  | -9.841% | 3.174% | 40.76% | 8.742% | -13.48% | -7.902% | 19.48% | 29.79% | -12.75% | -16.51% | -0.2679% | 10.76% | -3.055% | -0.6314% | -9.114% | -5.115% | -11.3% | 23.07% | -47.89% | 7.488% | -13.6% | -5.236% | 4.216% | -21.72%
| Z-Score (Absolute) |  | 0.66 | 0.27 | 2.98 | 0.67 | 0.92 | 0.52 | 1.45 | 2.19 | NA | 1.14 | 0.03 | 0.82 | 0.17 | 0 | NA | 0.32 | 0.77 | 1.7 | NA | 0.58 | 0.93 | 0.33 | 0.35 | 1.52
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C930: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1149: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1204: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 233b (Indicative Value 12.92 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.92 | 1
| 1 | 2 | 12.91 | 9
| Mean |  | 12.915 | 16
| Median |  | 12.915 | 12.915
| Std Dev. |  | 0.0070711 | 15
| Rel.Std.Dev. |  | 0.05475%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 233b (Indicative Value 9.17 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 9.16 | 1
| 1 | 2 | 9.18 | 10
| Mean |  | 9.17 | 16
| Median |  | 9.17 | 9.17
| Std Dev. |  | 0.014142 | 16
| Rel.Std.Dev. |  | 0.1542%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 233b (Indicative Value 11.65 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.65 | 1
| 1 | 2 | 11.64 | 11
| Mean |  | 11.645 | 16
| Median |  | 11.645 | 11.645
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.06072%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 233b (Indicative Value 0.628 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.627 | 1
| 1 | 2 | 0.628 | 12
| Mean |  | 0.6275 | 16
| Median |  | 0.6275 | 0.6275
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.1127%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 233b (Indicative Value 6.54 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.53 | 1
| 1 | 2 | 6.54 | 9
| Mean |  | 6.535 | 16
| Median |  | 6.535 | 6.535
| Std Dev. |  | 0.0070711 | 15
| Rel.Std.Dev. |  | 0.1082%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 233b (Indicative Value 0.179 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.179 | 1
| 1 | 2 | 0.179 | 10
| Mean |  | 0.179 | 16
| Median |  | 0.179 | 0.179
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 233b (Indicative Value 2.55 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.55 | 1
| 1 | 2 | 2.55 | 11
| Mean |  | 2.55 | 16
| Median |  | 2.55 | 2.55
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 233b (Indicative Value 0.094 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.095 | 1
| 1 | 2 | 0.093 | 12
| Mean |  | 0.094 | 16
| Median |  | 0.094 | 0.094
| Std Dev. |  | 0.0014142 | 18
| Rel.Std.Dev. |  | 1.504%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 233b (Indicative Value 0.309 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.3087 | 1
| 1 | 2 | 0.3091 | 16
| Mean |  | 0.3089 | 16
| Median |  | 0.3089 | 0.30895
| Std Dev. |  | 0.00028284 | 15
| Rel.Std.Dev. |  | 0.09156%
| PDM3 |  | -0.01527%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 233b (Indicative Value 51.83 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 51.79 | 1
| 1 | 2 | 51.86 | 10
| Mean |  | 51.825 | 16
| Median |  | 51.825 | 51.825
| Std Dev. |  | 0.049497 | 16
| Rel.Std.Dev. |  | 0.09551%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 233b (Indicative Value 1.04 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.041 | 1
| 1 | 2 | 1.034 | 11
| Mean |  | 1.0375 | 16
| Median |  | 1.0375 | 1.0375
| Std Dev. |  | 0.0049497 | 17
| Rel.Std.Dev. |  | 0.4771%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 233b (Indicative Value 3.18 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.19 | 1
| 1 | 2 | 3.16 | 14
| Mean |  | 3.175 | 16
| Median |  | 3.175 | 3.175
| Std Dev. |  | 0.021213 | 20
| Rel.Std.Dev. |  | 0.6681%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 233b (Indicative Value 0.13 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.13 | 1
| 1 | 2 | 0.13 | 16
| Mean |  | 0.13 | 16
| Median |  | 0.13 | 0.13
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 233b (Indicative Value 0.35 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.35 | 16
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 233b (Indicative Value 0.3 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.3 | 18
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 233b (Indicative Value 59 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 58.4 | 1
| 1 | 2 | 59.2 | 19
| Mean |  | 58.8 | 16
| Median |  | 58.8 | 58.8
| Std Dev. |  | 0.56569 | 25
| Rel.Std.Dev. |  | 0.9621%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 233b (Indicative Value 265 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 266 | 1
| 1 | 2 | 264 | 20
| Mean |  | 265 | 16
| Median |  | 265 | 265
| Std Dev. |  | 1.4142 | 26
| Rel.Std.Dev. |  | 0.5337%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 233b (Indicative Value 0.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 21
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 233b (Indicative Value 0.04 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.04 | 22
| Mean |  | 0.04 | 16
| Median |  | 0.04 | 0.04
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 233b (Indicative Value 0.6 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 23
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 233b (Indicative Value 12.3 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.3 | 1
| 1 | 2 | 12.2 | 24
| Mean |  | 12.25 | 16
| Median |  | 12.25 | 12.25
| Std Dev. |  | 0.070711 | 30
| Rel.Std.Dev. |  | 0.5772%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 233b (Indicative Value 44.8 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45 | 1
| 1 | 2 | 44.5 | 25
| Mean |  | 44.75 | 16
| Median |  | 44.75 | 44.75
| Std Dev. |  | 0.35355 | 31
| Rel.Std.Dev. |  | 0.7901%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 233b (Indicative Value 178 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 177 | 1
| 1 | 2 | 178 | 26
| Mean |  | 177.5 | 16
| Median |  | 177.5 | 177.5
| Std Dev. |  | 0.70711 | 32
| Rel.Std.Dev. |  | 0.3984%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 233b (Indicative Value 0.88 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.91 | 1
| 1 | 2 | 0.85 | 27
| Mean |  | 0.88 | 16
| Median |  | 0.88 | 0.88
| Std Dev. |  | 0.042426 | 33
| Rel.Std.Dev. |  | 4.821%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 233b (Indicative Value 166 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 168 | 28
| Mean |  | 166 | 16
| Median |  | 166 | 166
| Std Dev. |  | 2.8284 | 34
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 233b (Indicative Value 3.66 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.57 | 1
| 1 | 2 | 3.75 | 29
| Mean |  | 3.66 | 16
| Median |  | 3.66 | 3.66
| Std Dev. |  | 0.12728 | 35
| Rel.Std.Dev. |  | 3.478%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 233b (Indicative Value 2.26 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.23 | 1
| 1 | 2 | 2.28 | 30
| Mean |  | 2.255 | 16
| Median |  | 2.255 | 2.255
| Std Dev. |  | 0.035355 | 36
| Rel.Std.Dev. |  | 1.568%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 233b (Indicative Value 0.93 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.96 | 1
| 1 | 2 | 0.9 | 31
| Mean |  | 0.93 | 16
| Median |  | 0.93 | 0.93
| Std Dev. |  | 0.042426 | 37
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 233b (Indicative Value 15.2 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.2 | 1
| 1 | 2 | 15.2 | 32
| Mean |  | 15.2 | 16
| Median |  | 15.2 | 15.2
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 233b (Indicative Value 3.09 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.07 | 1
| 1 | 2 | 3.1 | 33
| Mean |  | 3.085 | 16
| Median |  | 3.085 | 3.085
| Std Dev. |  | 0.021213 | 39
| Rel.Std.Dev. |  | 0.6876%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 233b (Indicative Value 1.43 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.45 | 34
| Mean |  | 1.425 | 16
| Median |  | 1.425 | 1.425
| Std Dev. |  | 0.035355 | 40
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 233b (Indicative Value 1.88 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.92 | 1
| 1 | 2 | 1.83 | 18
| Mean |  | 1.875 | 16
| Median |  | 1.875 | 1.875
| Std Dev. |  | 0.06364 | 24
| Rel.Std.Dev. |  | 3.394%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 233b (Indicative Value 0.8 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.81 | 1
| 1 | 2 | 0.79 | 19
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0.014142 | 25
| Rel.Std.Dev. |  | 1.768%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 233b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 20
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 233b (Indicative Value 5.38 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.35 | 1
| 1 | 2 | 5.4 | 21
| Mean |  | 5.375 | 16
| Median |  | 5.375 | 5.375
| Std Dev. |  | 0.035355 | 27
| Rel.Std.Dev. |  | 0.6578%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 233b (Indicative Value 0.34 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.33 | 22
| Mean |  | 0.335 | 16
| Median |  | 0.335 | 0.335
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 2.111%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 233b (Indicative Value 0.157 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.156 | 1
| 1 | 2 | 0.158 | 23
| Mean |  | 0.157 | 16
| Median |  | 0.157 | 0.157
| Std Dev. |  | 0.0014142 | 29
| Rel.Std.Dev. |  | 0.9008%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 233b (Indicative Value 2.1 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2 | 24
| Mean |  | 2.1 | 16
| Median |  | 2.1 | 2.1
| Std Dev. |  | 0.14142 | 30
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 233b (Indicative Value 3.5 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.53 | 1
| 1 | 2 | 3.47 | 7
| Mean |  | 3.5 | 16
| Median |  | 3.5 | 3.5
| Std Dev. |  | 0.042426 | 31
| Rel.Std.Dev. |  | 1.212%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 233b (Indicative Value 8.55 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.73 | 1
| 1 | 2 | 8.36 | 26
| Mean |  | 8.545 | 16
| Median |  | 8.545 | 8.545
| Std Dev. |  | 0.26163 | 32
| Rel.Std.Dev. |  | 3.062%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 233b (Indicative Value 105 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 106 | 1
| 1 | 2 | 104 | 27
| Mean |  | 105 | 16
| Median |  | 105 | 105
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 233b (Indicative Value 24 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 24 | 1
| 1 | 2 | 24 | 28
| Mean |  | 24 | 16
| Median |  | 24 | 24
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 233b (Indicative Value 1.81 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.82 | 1
| 1 | 2 | 1.79 | 29
| Mean |  | 1.805 | 16
| Median |  | 1.805 | 1.805
| Std Dev. |  | 0.021213 | 35
| Rel.Std.Dev. |  | 1.175%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 233b (Indicative Value 13.6 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.6 | 1
| 1 | 2 | 13.6 | 30
| Mean |  | 13.6 | 16
| Median |  | 13.6 | 13.6
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 233b (Indicative Value 0.025 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.03 | 1
| 1 | 2 | 0.02 | 31
| Mean |  | 0.025 | 16
| Median |  | 0.025 | 0.025
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 233b (Indicative Value 1.3 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.3 | 32
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 233b (Indicative Value 38.9 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.7 | 1
| 1 | 2 | 39 | 33
| Mean |  | 38.85 | 16
| Median |  | 38.85 | 38.85
| Std Dev. |  | 0.21213 | 39
| Rel.Std.Dev. |  | 0.546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 233b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 7
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 40
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 233b (Indicative Value 2.47 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.49 | 1
| 1 | 2 | 2.44 | 18
| Mean |  | 2.465 | 16
| Median |  | 2.465 | 2.465
| Std Dev. |  | 0.035355 | 24
| Rel.Std.Dev. |  | 1.434%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 233b (Indicative Value 1 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1 | 19
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 233b (Indicative Value 192 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 192 | 1
| 1 | 2 | 191 | 20
| Mean |  | 191.5 | 16
| Median |  | 191.5 | 191.5
| Std Dev. |  | 0.70711 | 26
| Rel.Std.Dev. |  | 0.3692%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 233b (Indicative Value 0.2 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 21
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 233b (Indicative Value 0.55 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.55 | 1
| 1 | 2 | 0.54 | 22
| Mean |  | 0.545 | 16
| Median |  | 0.545 | 0.545
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 1.297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 233b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 29
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 233b (Indicative Value 0.99 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.98 | 1
| 1 | 2 | 1 | 24
| Mean |  | 0.99 | 16
| Median |  | 0.99 | 0.99
| Std Dev. |  | 0.014142 | 30
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 233b (Indicative Value 0.622 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.62 | 1
| 1 | 2 | 0.624 | 25
| Mean |  | 0.622 | 16
| Median |  | 0.622 | 0.622
| Std Dev. |  | 0.0028284 | 31
| Rel.Std.Dev. |  | 0.4547%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 233b (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 26
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 32
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 233b (Indicative Value 0.35 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.35 | 27
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 233b (Indicative Value 0.28 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.28 | 1
| 1 | 2 | 0.28 | 28
| Mean |  | 0.28 | 16
| Median |  | 0.28 | 0.28
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 233b (Indicative Value 275 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 275 | 1
| 1 | 2 | 274 | 29
| Mean |  | 274.5 | 16
| Median |  | 274.5 | 274.5
| Std Dev. |  | 0.70711 | 35
| Rel.Std.Dev. |  | 0.2576%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 233b (Indicative Value 30 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 30
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 233b (Indicative Value 21.2 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.7 | 1
| 1 | 2 | 20.7 | 31
| Mean |  | 21.2 | 16
| Median |  | 21.2 | 21.2
| Std Dev. |  | 0.70711 | 37
| Rel.Std.Dev. |  | 3.335%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 233b (Indicative Value 2.25 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.29 | 32
| Mean |  | 2.245 | 16
| Median |  | 2.245 | 2.245
| Std Dev. |  | 0.06364 | 38
| Rel.Std.Dev. |  | 2.835%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 233b (Indicative Value 145 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 145 | 1
| 1 | 2 | 145 | 33
| Mean |  | 145 | 16
| Median |  | 145 | 145
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 233b (Indicative Value 65 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 65 | 1
| 1 | 2 | 65 | 34
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.