## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 234b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 1.2342 | 1.2205 | 1.2479 | 1.2251 | 1.2433
PhotonAssay (recommended gross mass* 355±15 g)
Au, Gold (ppm) | 1.2195 | 1.2109 | 1.2281 | 1.217 | 1.2219
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 1.1724 | 1.1531 | 1.1918 | 1.163 | 1.1819
Cyanide Leach
Au, Gold (ppm) | 1.1731 | 1.1463 | 1.1998 | 1.1697 | 1.1764
4-Acid Digestion
Ag, Silver (ppm) | 0.34803 | 0.325 | 0.37106 | 0.32208 | 0.37397
Al, Aluminium (wt.%) | 6.7183 | 6.5393 | 6.8974 | 6.5782 | 6.8584
As, Arsenic (ppm) | 65.2859 | 63.1065 | 67.4652 | 63.009 | 67.5627
Ba, Barium (ppm) | 289.8811 | 277.9853 | 301.7768 | 277.4287 | 302.3334
Be, Beryllium (ppm) | 0.43576 | 0.40231 | 0.46921 | 0.41545 | 0.45608
Bi, Bismuth (ppm) | 0.047385 | 0.036519 | 0.05825 | IND | IND
Ca, Calcium (wt.%) | 6.146 | 5.9838 | 6.3081 | 6.0461 | 6.2459
Cd, Cadmium (ppm) | 0.61798 | 0.56985 | 0.6661 | 0.58605 | 0.6499
Ce, Cerium (ppm) | 13.0911 | 12.5004 | 13.6818 | 12.5937 | 13.5884
Co, Cobalt (ppm) | 43.195 | 41.6954 | 44.6946 | 41.8678 | 44.5222
Cr, Chromium (ppm) | 144.1065 | 135.5898 | 152.6231 | 138.8404 | 149.3726
Cs, Caesium (ppm) | 0.95225 | 0.90778 | 0.99671 | 0.91196 | 0.99253
Cu, Copper (ppm) | 161.8509 | 156.8198 | 166.8819 | 158.5106 | 165.1912
Dy, Dysprosium (ppm) | 3.5898 | 3.3464 | 3.8332 | 3.5068 | 3.6728
Er, Erbium (ppm) | 2.1733 | 2.0122 | 2.3344 | 2.0761 | 2.2706
Eu, Europium (ppm) | 0.91333 | 0.84182 | 0.98484 | 0.87061 | 0.95604
Fe, Iron (wt.%) | 7.8271 | 7.6209 | 8.0333 | 7.6742 | 7.9801
Ga, Gallium (ppm) | 15.239 | 14.5547 | 15.9234 | 14.7113 | 15.7668
Gd, Gadolinium (ppm) | 2.9638 | 2.6976 | 3.2299 | 2.8675 | 3.06
Ge, Germanium (ppm) | 0.087583 | 0.067905 | 0.10726 | IND | IND
Hf, Hafnium (ppm) | 1.6435 | 1.4928 | 1.7943 | 1.546 | 1.7411
Ho, Holmium (ppm) | 0.73643 | 0.67168 | 0.80119 | 0.70623 | 0.76664
In, Indium (ppm) | 0.074811 | 0.065363 | 0.084258 | 0.066596 | 0.083026
K, Potassium (wt.%) | 0.57775 | 0.55918 | 0.59632 | 0.56052 | 0.59497
La, Lanthanum (ppm) | 5.6128 | 5.3385 | 5.8871 | 5.3802 | 5.8454
Li, Lithium (ppm) | 12.1613 | 11.5164 | 12.8063 | 11.7767 | 12.546
Lu, Lutetium (ppm) | 0.31269 | 0.27403 | 0.35135 | 0.29965 | 0.32573
Mg, Magnesium (wt.%) | 3.7644 | 3.6653 | 3.8636 | 3.6882 | 3.8407
Mn, Manganese (wt.%) | 0.14785 | 0.14363 | 0.15208 | 0.14501 | 0.1507
Mo, Molybdenum (ppm) | 2.5666 | 2.4081 | 2.725 | 2.4602 | 2.673
Na, Sodium (wt.%) | 1.942 | 1.8817 | 2.0022 | 1.8879 | 1.996
Nb, Niobium (ppm) | 3.5911 | 3.3281 | 3.8542 | 3.3986 | 3.7837
Nd, Neodymium (ppm) | 8.3934 | 7.9172 | 8.8696 | 8.1436 | 8.6432
Ni, Nickel (ppm) | 96.9143 | 93.8077 | 100.0208 | 94.786 | 99.0425
P, Phosphorus (wt.%) | 0.043213 | 0.041661 | 0.044764 | 0.042257 | 0.044168
Pb, Lead (ppm) | 26.2337 | 25.0956 | 27.3717 | 25.269 | 27.1983
Pr, Praseodymium (ppm) | 1.7507 | 1.6436 | 1.8577 | 1.6682 | 1.8331
Rb, Rubidium (ppm) | 15.1797 | 14.4438 | 15.9157 | 14.61 | 15.7495
S, Sulphur (wt.%) | 0.40505 | 0.3934 | 0.41671 | 0.39482 | 0.41529
Sb, Antimony (ppm) | 1.4269 | 1.3247 | 1.5292 | 1.3188 | 1.5351
Sc, Scandium (ppm) | 36.7547 | 35.2233 | 38.2861 | 35.6358 | 37.8736
Sm, Samarium (ppm) | 2.4849 | 2.2867 | 2.6832 | 2.3524 | 2.6174
Sn, Tin (ppm) | 1.1388 | 1.0289 | 1.2487 | 1.0229 | 1.2547
Sr, Strontium (ppm) | 193.6714 | 185.7403 | 201.6024 | 188.532 | 198.8108
Ta, Tantalum (ppm) | 0.2386 | 0.20923 | 0.26797 | 0.21906 | 0.25814
Tb, Terbium (ppm) | 0.54054 | 0.49686 | 0.58421 | 0.51241 | 0.56866
Te, Tellurium (ppm) | 0.10103 | 0.069027 | 0.13303 | IND | IND
Th, Thorium (ppm) | 1.0159 | 0.94548 | 1.0863 | 0.97375 | 1.058
Ti, Titanium (wt.%) | 0.59177 | 0.57475 | 0.60878 | 0.57872 | 0.60481
Tl, Thallium (ppm) | 0.2266 | 0.21485 | 0.23836 | 0.21377 | 0.23943
Tm, Thulium (ppm) | 0.31326 | 0.28132 | 0.34521 | 0.29498 | 0.33155
U, Uranium (ppm) | 0.30383 | 0.29326 | 0.31441 | 0.2891 | 0.31857
V, Vanadium (ppm) | 257.6248 | 249.6736 | 265.5761 | 251.5515 | 263.6982
W, Tungsten (ppm) | 33.305 | 31.7381 | 34.8719 | 32.3484 | 34.2617
Y, Yttrium (ppm) | 19.4404 | 18.5634 | 20.3173 | 18.9796 | 19.9012
Yb, Ytterbium (ppm) | 2.0336 | 1.8633 | 2.2039 | 1.942 | 2.1252
Zn, Zinc (ppm) | 143.0977 | 139.0096 | 147.1857 | 139.9082 | 146.2872
Zr, Zirconium (ppm) | 49.5548 | 45.9015 | 53.2081 | 47.5973 | 51.5123
Aqua Regia Digestion
Ag, Silver (ppm) | 0.34318 | 0.31614 | 0.37023 | 0.32279 | 0.36357
Al, Aluminium (wt.%) | 3.1165 | 2.9936 | 3.2394 | 3.047 | 3.186
As, Arsenic (ppm) | 64.9188 | 62.8013 | 67.0362 | 63.2723 | 66.5652
B, Boron (ppm) | 26.851 | 21.2811 | 32.421 | 25.3265 | 28.3756
Ba, Barium (ppm) | 41.1843 | 38.9247 | 43.4439 | 39.8904 | 42.4781
Be, Beryllium (ppm) | 0.21139 | 0.18542 | 0.23736 | 0.20189 | 0.22088
Bi, Bismuth (ppm) | 0.042285 | 0.032633 | 0.051936 | IND | IND
Ca, Calcium (wt.%) | 2.201 | 2.0571 | 2.3448 | 2.1425 | 2.2594
Cd, Cadmium (ppm) | 0.59468 | 0.56006 | 0.62929 | 0.56928 | 0.62007
Ce, Cerium (ppm) | 9.7346 | 9.3582 | 10.1109 | 9.4367 | 10.0324
Co, Cobalt (ppm) | 31.4269 | 30.3109 | 32.5429 | 30.7753 | 32.0784
Cr, Chromium (ppm) | 30.041 | 28.5051 | 31.5769 | 28.9761 | 31.1059
Cs, Caesium (ppm) | 0.71487 | 0.68796 | 0.74178 | 0.69313 | 0.7366
Cu, Copper (ppm) | 160.3635 | 156.2083 | 164.5188 | 157.1308 | 163.5962
Dy, Dysprosium (ppm) | 2.0461 | 1.6615 | 2.4307 | 1.9603 | 2.1319
Er, Erbium (ppm) | 1.1783 | 0.96239 | 1.3942 | 1.1225 | 1.2341
Eu, Europium (ppm) | 0.44925 | 0.34208 | 0.55641 | 0.41685 | 0.48164
Fe, Iron (wt.%) | 5.7077 | 5.5333 | 5.8821 | 5.6116 | 5.8037
Ga, Gallium (ppm) | 9.7786 | 9.2002 | 10.357 | 9.5621 | 9.9951
Gd, Gadolinium (ppm) | 1.8084 | 1.5548 | 2.062 | 1.7537 | 1.8631
Ge, Germanium (ppm) | 0.104 | 0.077316 | 0.13068 | IND | IND
Hf, Hafnium (ppm) | 0.4091 | 0.35657 | 0.46162 | 0.38141 | 0.43678
Ho, Holmium (ppm) | 0.39221 | 0.32026 | 0.46416 | 0.36603 | 0.41838
In, Indium (ppm) | 0.030514 | 0.027634 | 0.033394 | 0.027406 | 0.033622
K, Potassium (wt.%) | 0.15082 | 0.14425 | 0.15739 | 0.14724 | 0.1544
La, Lanthanum (ppm) | 4.3818 | 4.2061 | 4.5575 | 4.212 | 4.5517
Li, Lithium (ppm) | 10.3401 | 9.8097 | 10.8706 | 10.0421 | 10.6381
Lu, Lutetium (ppm) | 0.1337 | 0.11769 | 0.14971 | IND | IND
Mg, Magnesium (wt.%) | 1.8022 | 1.754 | 1.8504 | 1.7717 | 1.8327
Mn, Manganese (wt.%) | 0.081662 | 0.078889 | 0.084435 | 0.080246 | 0.083078
Mo, Molybdenum (ppm) | 2.4136 | 2.3119 | 2.5153 | 2.324 | 2.5033
Na, Sodium (wt.%) | 0.2405 | 0.23041 | 0.25058 | 0.23015 | 0.25084
Nd, Neodymium (ppm) | 5.7785 | 5.3036 | 6.2533 | 5.633 | 5.9239
Ni, Nickel (ppm) | 67.8733 | 65.4342 | 70.3124 | 65.9821 | 69.7645
P, Phosphorus (wt.%) | 0.042484 | 0.041115 | 0.043853 | 0.041519 | 0.043449
Pb, Lead (ppm) | 25.8888 | 24.9777 | 26.7999 | 25.2156 | 26.562
Pr, Praseodymium (ppm) | 1.2274 | 1.1081 | 1.3466 | 1.1847 | 1.27
Rb, Rubidium (ppm) | 6.3303 | 6.0473 | 6.6132 | 6.1761 | 6.4844
S, Sulphur (wt.%) | 0.40163 | 0.38967 | 0.41358 | 0.39131 | 0.41194
Sb, Antimony (ppm) | 0.8721 | 0.75639 | 0.98781 | 0.82995 | 0.91425
Sc, Scandium (ppm) | 4.9578 | 4.5326 | 5.3829 | 4.7869 | 5.1286
Sm, Samarium (ppm) | 1.5174 | 1.3699 | 1.6648 | 1.4615 | 1.5733
Sn, Tin (ppm) | 0.65094 | 0.5917 | 0.71018 | 0.59416 | 0.70771
Sr, Strontium (ppm) | 36.9431 | 33.8113 | 40.075 | 35.4758 | 38.4104
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.30411 | 0.28033 | 0.32789 | 0.28941 | 0.31882
Te, Tellurium (ppm) | 0.10124 | 0.076415 | 0.12607 | IND | IND
Th, Thorium (ppm) | 0.80163 | 0.75093 | 0.85233 | 0.77837 | 0.82489
Ti, Titanium (wt.%) | 0.31453 | 0.28641 | 0.34266 | 0.30325 | 0.32582
Tl, Thallium (ppm) | 0.11831 | 0.11059 | 0.12602 | 0.097624 | 0.13899
Tm, Thulium (ppm) | 0.15719 | 0.10159 | 0.21279 | IND | IND
U, Uranium (ppm) | 0.20321 | 0.19034 | 0.21608 | 0.19061 | 0.21581
V, Vanadium (ppm) | 127.4601 | 119.2483 | 135.6719 | 123.8791 | 131.0411
W, Tungsten (ppm) | 23.5262 | 21.6351 | 25.4172 | 22.7297 | 24.3227
Y, Yttrium (ppm) | 10.6429 | 10.3095 | 10.9764 | 10.4011 | 10.8848
Yb, Ytterbium (ppm) | 0.97326 | 0.81462 | 1.1319 | 0.903 | 1.0435
Zn, Zinc (ppm) | 129.4807 | 125.2544 | 133.7071 | 127.0901 | 131.8714
Zr, Zirconium (ppm) | 13.9286 | 12.7427 | 15.1145 | 13.3603 | 14.4969
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 234b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 18.5625 | Pt | ppb | 10.0121
4-Acid Digestion
B | ppm | 47.6333 | Re | ppm | 0.0030606
Hg | ppm | 0.047033 | Se | ppm | 0.98633
Aqua Regia Digestion
Hg | ppm | 0.048114 | Pd | ppb | 16.0556 | Se | ppm | 0.65583
Ir | ppm | < 0.003 | Pt | ppb | 7.7663 | Si | wt.% | 0.092962
Nb | ppm | 0.12197 | Re | ppm | 0.0025606
Borate Fusion XRF
Al2O3 | wt.% | 12.81 | MgO | wt.% | 6.43 | S | wt.% | 0.33337
CaO | wt.% | 8.94 | MnO | wt.% | 0.18 | SiO2 | wt.% | 52.19
Fe2O3 | wt.% | 11.615 | Na2O | wt.% | 2.605 | TiO2 | wt.% | 1.0365
K2O | wt.% | 0.673 | P2O5 | wt.% | 0.095
Thermogravimetry
LOI1000 | wt.% | 3.13
Infrared Combustion
C | wt.% | 0.14 | S | wt.% | 0.375
Laser Ablation ICP-MS
Ag | ppm | 0.35 | Hf | ppm | 1.87 | Sm | ppm | 2.475
As | ppm | 63.3 | Ho | ppm | 0.8 | Sn | ppm | 1
Ba | ppm | 286.5 | In | ppm | 0.05 | Sr | ppm | 188
Be | ppm | 0.6 | La | ppm | 5.63 | Ta | ppm | 0.235
Bi | ppm | 0.06 | Lu | ppm | 0.325 | Tb | ppm | 0.565
Cd | ppm | 0.7 | Mn | wt.% | 0.1565 | Te | ppm | < 0.2
Ce | ppm | 12.55 | Mo | ppm | 2.4 | Th | ppm | 1.01
Co | ppm | 43.9 | Nb | ppm | 3.615 | Ti | wt.% | 0.618
Cr | ppm | 167 | Nd | ppm | 8.525 | Tl | ppm | < 0.2
Cs | ppm | 0.91 | Ni | ppm | 102 | Tm | ppm | 0.35
Cu | ppm | 167 | Pb | ppm | 26.5 | U | ppm | 0.31
Dy | ppm | 3.6 | Pr | ppm | 1.845 | V | ppm | 268
Er | ppm | 2.37 | Rb | ppm | 14.5 | W | ppm | 32.75
Eu | ppm | 0.885 | Re | ppm | 0.0125 | Y | ppm | 20.1
Ga | ppm | 14.65 | Sb | ppm | 1.5 | Yb | ppm | 2.25
Gd | ppm | 3.12 | Sc | ppm | 38 | Zn | ppm | 172.5
Ge | ppm | 1.425 | Se | ppm | < 5 | Zr | ppm | 65.25
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 234b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 1.2342 | 0.036823 | 1.1606 | 1.3079 | 1.1238 | 1.3447 | 2.983% | 5.967% | 8.95% | 1.1725 | 1.2959
PhotonAssay
Au, ppm | 1.2195 | 0.050192 | 1.1191 | 1.3199 | 1.0689 | 1.37 | 4.116% | 8.232% | 12.35% | 1.1585 | 1.2804
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.1724 | 0.045234 | 1.082 | 1.2629 | 1.0367 | 1.3081 | 3.858% | 7.716% | 11.57% | 1.1138 | 1.2311
Cyanide Leach
Au, ppm | 1.1731 | 0.058419 | 1.0562 | 1.2899 | 0.99781 | 1.3483 | 4.98% | 9.96% | 14.94% | 1.1144 | 1.2317
4-Acid Digestion
Ag, ppm | 0.34803 | 0.020991 | 0.30605 | 0.39001 | 0.28505 | 0.411 | 6.031% | 12.06% | 18.09% | 0.33063 | 0.36543
Al, wt.% | 6.7183 | 0.22319 | 6.2719 | 7.1647 | 6.0487 | 7.3879 | 3.322% | 6.644% | 9.966% | 6.3824 | 7.0542
As, ppm | 65.2859 | 2.9211 | 59.4437 | 71.1281 | 56.5226 | 74.0492 | 4.474% | 8.949% | 13.42% | 62.0216 | 68.5502
Ba, ppm | 289.8811 | 10.0834 | 269.7143 | 310.0478 | 259.6309 | 320.1312 | 3.478% | 6.957% | 10.44% | 275.387 | 304.3751
Be, ppm | 0.43576 | 0.05029 | 0.33518 | 0.53634 | 0.28489 | 0.58663 | 11.54% | 23.08% | 34.62% | 0.41398 | 0.45755
Bi, ppm | 0.047385 | 0.0068126 | 0.033759 | 0.06101 | 0.026947 | 0.067822 | 14.38% | 28.75% | 43.13% | 0.045015 | 0.049754
Ca, wt.% | 6.146 | 0.24781 | 5.6504 | 6.6416 | 5.4025 | 6.8894 | 4.032% | 8.064% | 12.1% | 5.8387 | 6.4533
Cd, ppm | 0.61798 | 0.040749 | 0.53648 | 0.69947 | 0.49573 | 0.74022 | 6.594% | 13.19% | 19.78% | 0.58708 | 0.64887
Ce, ppm | 13.0911 | 0.80611 | 11.4789 | 14.7033 | 10.6728 | 15.5094 | 6.158% | 12.32% | 18.47% | 12.4365 | 13.7456
Co, ppm | 43.195 | 2.4878 | 38.2194 | 48.1707 | 35.7316 | 50.6585 | 5.76% | 11.52% | 17.28% | 41.0353 | 45.3548
Cr, ppm | 144.1065 | 16.6383 | 110.8298 | 177.3831 | 94.1915 | 194.0215 | 11.55% | 23.09% | 34.64% | 136.9011 | 151.3118
Cs, ppm | 0.95225 | 0.049677 | 0.85289 | 1.0516 | 0.80321 | 1.1013 | 5.217% | 10.43% | 15.65% | 0.90463 | 0.99986
Cu, ppm | 161.8509 | 7.4214 | 147.0082 | 176.6936 | 139.5868 | 184.1149 | 4.585% | 9.171% | 13.76% | 153.7583 | 169.9434
Dy, ppm | 3.5898 | 0.16241 | 3.265 | 3.9146 | 3.1026 | 4.077 | 4.524% | 9.048% | 13.57% | 3.4103 | 3.7693
Er, ppm | 2.1733 | 0.12888 | 1.9156 | 2.4311 | 1.7867 | 2.56 | 5.93% | 11.86% | 17.79% | 2.0647 | 2.282
Eu, ppm | 0.91333 | 0.050521 | 0.81228 | 1.0144 | 0.76176 | 1.0649 | 5.532% | 11.06% | 16.59% | 0.86766 | 0.95899
Fe, wt.% | 7.8271 | 0.33027 | 7.1666 | 8.4877 | 6.8363 | 8.8179 | 4.22% | 8.439% | 12.66% | 7.4358 | 8.2185
Ga, ppm | 15.239 | 0.70474 | 13.8295 | 16.6485 | 13.1248 | 17.3533 | 4.625% | 9.249% | 13.87% | 14.4771 | 16.001
Gd, ppm | 2.9638 | 0.2108 | 2.5422 | 3.3854 | 2.3313 | 3.5962 | 7.113% | 14.23% | 21.34% | 2.8156 | 3.1119
Ge, ppm | 0.087583 | 0.020068 | 0.047447 | 0.12772 | 0.027379 | 0.14779 | 22.91% | 45.83% | 68.74% | 0.083204 | 0.091962
Hf, ppm | 1.6435 | 0.18548 | 1.2726 | 2.0145 | 1.0871 | 2.2 | 11.29% | 22.57% | 33.86% | 1.5613 | 1.7257
Ho, ppm | 0.73643 | 0.044571 | 0.64729 | 0.82557 | 0.60272 | 0.87015 | 6.052% | 12.1% | 18.16% | 0.69961 | 0.77325
In, ppm | 0.074811 | 0.0072171 | 0.060376 | 0.089245 | 0.053159 | 0.096462 | 9.647% | 19.29% | 28.94% | 0.07107 | 0.078551
K, wt.% | 0.57775 | 0.026415 | 0.52492 | 0.63058 | 0.4985 | 0.65699 | 4.572% | 9.144% | 13.72% | 0.54886 | 0.60663
La, ppm | 5.6128 | 0.39258 | 4.8276 | 6.398 | 4.4351 | 6.7905 | 6.994% | 13.99% | 20.98% | 5.3322 | 5.8934
Li, ppm | 12.1613 | 0.84341 | 10.4745 | 13.8482 | 9.6311 | 14.6916 | 6.935% | 13.87% | 20.81% | 11.5533 | 12.7694
Lu, ppm | 0.31269 | 0.025952 | 0.26079 | 0.3646 | 0.23484 | 0.39055 | 8.3% | 16.6% | 24.9% | 0.29706 | 0.32833
Mg, wt.% | 3.7644 | 0.16316 | 3.4381 | 4.0907 | 3.275 | 4.2539 | 4.334% | 8.668% | 13% | 3.5762 | 3.9526
Mn, wt.% | 0.14785 | 0.0046766 | 0.1385 | 0.15721 | 0.13382 | 0.16188 | 3.163% | 6.326% | 9.489% | 0.14046 | 0.15525
Mo, ppm | 2.5666 | 0.23942 | 2.0877 | 3.0454 | 1.8483 | 3.2849 | 9.329% | 18.66% | 27.99% | 2.4383 | 2.6949
Na, wt.% | 1.942 | 0.072771 | 1.7964 | 2.0875 | 1.7236 | 2.1603 | 3.747% | 7.495% | 11.24% | 1.8449 | 2.0391
Nb, ppm | 3.5911 | 0.29026 | 3.0106 | 4.1717 | 2.7204 | 4.4619 | 8.083% | 16.17% | 24.25% | 3.4116 | 3.7707
Nd, ppm | 8.3934 | 0.34268 | 7.708 | 9.0788 | 7.3654 | 9.4214 | 4.083% | 8.165% | 12.25% | 7.9737 | 8.8131
Ni, ppm | 96.9143 | 3.9972 | 88.9198 | 104.9087 | 84.9226 | 108.9059 | 4.124% | 8.249% | 12.37% | 92.0686 | 101.76
P, wt.% | 0.043213 | 0.0013414 | 0.04053 | 0.045895 | 0.039188 | 0.047237 | 3.104% | 6.208% | 9.313% | 0.041052 | 0.045373
Pb, ppm | 26.2337 | 1.3593 | 23.5151 | 28.9522 | 22.1559 | 30.3114 | 5.181% | 10.36% | 15.54% | 24.922 | 27.5453
Pr, ppm | 1.7507 | 0.081138 | 1.5884 | 1.9129 | 1.5072 | 1.9941 | 4.635% | 9.269% | 13.9% | 1.6631 | 1.8382
Rb, ppm | 15.1797 | 0.91393 | 13.3519 | 17.0076 | 12.4379 | 17.9215 | 6.021% | 12.04% | 18.06% | 14.4208 | 15.9387
S, wt.% | 0.40505 | 0.0099915 | 0.38507 | 0.42504 | 0.37508 | 0.43503 | 2.467% | 4.933% | 7.4% | 0.3848 | 0.42531
Sb, ppm | 1.4269 | 0.08618 | 1.2546 | 1.5993 | 1.1684 | 1.6855 | 6.039% | 12.08% | 18.12% | 1.3556 | 1.4983
Sc, ppm | 36.7547 | 2.2239 | 32.307 | 41.2025 | 30.0831 | 43.4264 | 6.051% | 12.1% | 18.15% | 34.917 | 38.5925
Sm, ppm | 2.4849 | 0.10995 | 2.265 | 2.7048 | 2.1551 | 2.8148 | 4.425% | 8.85% | 13.27% | 2.3607 | 2.6092
Sn, ppm | 1.1388 | 0.1685 | 0.80179 | 1.4758 | 0.63329 | 1.6443 | 14.8% | 29.59% | 44.39% | 1.0819 | 1.1957
Sr, ppm | 193.6714 | 9.085 | 175.5013 | 211.8414 | 166.4163 | 220.9265 | 4.691% | 9.382% | 14.07% | 183.9878 | 203.3549
Ta, ppm | 0.2386 | 0.0421 | 0.1544 | 0.3228 | 0.1123 | 0.3649 | 17.64% | 35.29% | 52.93% | 0.22667 | 0.25053
Tb, ppm | 0.54054 | 0.04329 | 0.45396 | 0.62712 | 0.41067 | 0.67041 | 8.009% | 16.02% | 24.03% | 0.51351 | 0.56756
Te, ppm | 0.10103 | 0.021686 | 0.057658 | 0.1444 | 0.035972 | 0.16609 | 21.46% | 42.93% | 64.39% | 0.095979 | 0.10608
Th, ppm | 1.0159 | 0.073865 | 0.86816 | 1.1636 | 0.79429 | 1.2375 | 7.271% | 14.54% | 21.81% | 0.96509 | 1.0667
Ti, wt.% | 0.59177 | 0.014002 | 0.56376 | 0.61977 | 0.54976 | 0.63377 | 2.366% | 4.732% | 7.098% | 0.56218 | 0.62135
Tl, ppm | 0.2266 | 0.014045 | 0.19851 | 0.25469 | 0.18447 | 0.26874 | 6.198% | 12.4% | 18.59% | 0.21527 | 0.23793
Tm, ppm | 0.31326 | 0.023841 | 0.26558 | 0.36095 | 0.24174 | 0.38479 | 7.611% | 15.22% | 22.83% | 0.2976 | 0.32893
U, ppm | 0.30383 | 0.0099824 | 0.28387 | 0.3238 | 0.27389 | 0.33378 | 3.285% | 6.571% | 9.856% | 0.28864 | 0.31903
V, ppm | 257.6248 | 9.8478 | 237.9292 | 277.3205 | 228.0813 | 287.1683 | 3.823% | 7.645% | 11.47% | 244.7436 | 270.5061
W, ppm | 33.305 | 2.565 | 28.1751 | 38.435 | 25.6101 | 40.9999 | 7.701% | 15.4% | 23.1% | 31.6398 | 34.9703
Y, ppm | 19.4404 | 1.1881 | 17.0641 | 21.8167 | 15.8759 | 23.0048 | 6.112% | 12.22% | 18.34% | 18.4684 | 20.4124
Yb, ppm | 2.0336 | 0.16817 | 1.6973 | 2.3699 | 1.5291 | 2.5381 | 8.27% | 16.54% | 24.81% | 1.9319 | 2.1353
Zn, ppm | 143.0977 | 7.1376 | 128.8225 | 157.3729 | 121.6849 | 164.5105 | 4.988% | 9.976% | 14.96% | 135.9428 | 150.2526
Zr, ppm | 49.5548 | 6.3011 | 36.9526 | 62.1569 | 30.6516 | 68.458 | 12.72% | 25.43% | 38.15% | 47.0771 | 52.0325
Aqua Regia Digestion
Ag, ppm | 0.34318 | 0.021274 | 0.30063 | 0.38573 | 0.27936 | 0.40701 | 6.199% | 12.4% | 18.6% | 0.32602 | 0.36034
Al, wt.% | 3.1165 | 0.25852 | 2.5995 | 3.6336 | 2.3409 | 3.8921 | 8.295% | 16.59% | 24.89% | 2.9607 | 3.2723
As, ppm | 64.9188 | 2.8256 | 59.2675 | 70.57 | 56.4418 | 73.3957 | 4.353% | 8.705% | 13.06% | 61.6728 | 68.1647
B, ppm | 26.851 | 7.2722 | 12.3067 | 41.3953 | 5.0346 | 48.6675 | 27.08% | 54.17% | 81.25% | 25.5085 | 28.1936
Ba, ppm | 41.1843 | 4.3195 | 32.5453 | 49.8233 | 28.2257 | 54.1429 | 10.49% | 20.98% | 31.46% | 39.1251 | 43.2435
Be, ppm | 0.21139 | 0.035245 | 0.1409 | 0.28188 | 0.10565 | 0.31712 | 16.67% | 33.35% | 50.02% | 0.20082 | 0.22196
Bi, ppm | 0.042285 | 0.010394 | 0.021497 | 0.063072 | 0.011104 | 0.073465 | 24.58% | 49.16% | 73.74% | 0.04017 | 0.044399
Ca, wt.% | 2.201 | 0.33298 | 1.535 | 2.8669 | 1.202 | 3.1999 | 15.13% | 30.26% | 45.39% | 2.0909 | 2.311
Cd, ppm | 0.59468 | 0.046074 | 0.50253 | 0.68682 | 0.45645 | 0.7329 | 7.748% | 15.5% | 23.24% | 0.56494 | 0.62441
Ce, ppm | 9.7346 | 0.42191 | 8.8908 | 10.5784 | 8.4689 | 11.0003 | 4.334% | 8.668% | 13% | 9.2478 | 10.2213
Co, ppm | 31.4269 | 2.0649 | 27.2971 | 35.5566 | 25.2322 | 37.6215 | 6.57% | 13.14% | 19.71% | 29.8555 | 32.9982
Cr, ppm | 30.041 | 2.7922 | 24.4566 | 35.6254 | 21.6644 | 38.4176 | 9.295% | 18.59% | 27.88% | 28.5389 | 31.543
Cs, ppm | 0.71487 | 0.038721 | 0.63742 | 0.79231 | 0.5987 | 0.83103 | 5.417% | 10.83% | 16.25% | 0.67912 | 0.75061
Cu, ppm | 160.3635 | 5.8598 | 148.6439 | 172.0832 | 142.7841 | 177.943 | 3.654% | 7.308% | 10.96% | 152.3454 | 168.3817
Dy, ppm | 2.0461 | 0.27992 | 1.4863 | 2.606 | 1.2063 | 2.8859 | 13.68% | 27.36% | 41.04% | 1.9438 | 2.1484
Er, ppm | 1.1783 | 0.17046 | 0.83737 | 1.5192 | 0.66691 | 1.6897 | 14.47% | 28.93% | 43.4% | 1.1194 | 1.2372
Eu, ppm | 0.44925 | 0.082021 | 0.2852 | 0.61329 | 0.20318 | 0.69531 | 18.26% | 36.52% | 54.77% | 0.42678 | 0.47171
Fe, wt.% | 5.7077 | 0.22876 | 5.2501 | 6.1652 | 5.0214 | 6.3939 | 4.008% | 8.016% | 12.02% | 5.4223 | 5.993
Ga, ppm | 9.7786 | 0.92237 | 7.9338 | 11.6233 | 7.0115 | 12.5457 | 9.433% | 18.87% | 28.3% | 9.2897 | 10.2675
Gd, ppm | 1.8084 | 0.18328 | 1.4418 | 2.175 | 1.2586 | 2.3582 | 10.13% | 20.27% | 30.4% | 1.718 | 1.8988
Ge, ppm | 0.104 | 0.027935 | 0.048131 | 0.15987 | 0.020196 | 0.1878 | 26.86% | 53.72% | 80.58% | 0.0988 | 0.1092
Hf, ppm | 0.4091 | 0.048008 | 0.31308 | 0.50511 | 0.26507 | 0.55312 | 11.74% | 23.47% | 35.21% | 0.38864 | 0.42955
Ho, ppm | 0.39221 | 0.037967 | 0.31627 | 0.46814 | 0.27831 | 0.50611 | 9.68% | 19.36% | 29.04% | 0.3726 | 0.41182
In, ppm | 0.030514 | 0.0020931 | 0.026328 | 0.0347 | 0.024235 | 0.036793 | 6.86% | 13.72% | 20.58% | 0.028988 | 0.03204
K, wt.% | 0.15082 | 0.0098711 | 0.13108 | 0.17056 | 0.12121 | 0.18043 | 6.545% | 13.09% | 19.63% | 0.14328 | 0.15836
La, ppm | 4.3818 | 0.29436 | 3.7931 | 4.9705 | 3.4988 | 5.2649 | 6.718% | 13.44% | 20.15% | 4.1627 | 4.6009
Li, ppm | 10.3401 | 0.77803 | 8.7841 | 11.8962 | 8.006 | 12.6742 | 7.524% | 15.05% | 22.57% | 9.8231 | 10.8571
Lu, ppm | 0.1337 | 0.022977 | 0.087746 | 0.17965 | 0.064769 | 0.20263 | 17.19% | 34.37% | 51.56% | 0.12701 | 0.14038
Mg, wt.% | 1.8022 | 0.090688 | 1.6208 | 1.9836 | 1.5301 | 2.0742 | 5.032% | 10.06% | 15.1% | 1.7121 | 1.8923
Mn, wt.% | 0.081662 | 0.0057778 | 0.070107 | 0.093218 | 0.064329 | 0.098995 | 7.075% | 14.15% | 21.23% | 0.077579 | 0.085745
Mo, ppm | 2.4136 | 0.10301 | 2.2076 | 2.6196 | 2.1046 | 2.7227 | 4.268% | 8.535% | 12.8% | 2.293 | 2.5343
Na, wt.% | 0.2405 | 0.017463 | 0.20557 | 0.27542 | 0.18811 | 0.29289 | 7.261% | 14.52% | 21.78% | 0.22847 | 0.25252
Nd, ppm | 5.7785 | 0.44102 | 4.8964 | 6.6605 | 4.4554 | 7.1015 | 7.632% | 15.26% | 22.9% | 5.4895 | 6.0674
Ni, ppm | 67.8733 | 3.4187 | 61.0359 | 74.7107 | 57.6172 | 78.1294 | 5.037% | 10.07% | 15.11% | 64.4796 | 71.267
P, wt.% | 0.042484 | 0.0020315 | 0.038421 | 0.046547 | 0.036389 | 0.048578 | 4.782% | 9.564% | 14.35% | 0.04036 | 0.044608
Pb, ppm | 25.8888 | 1.2355 | 23.4177 | 28.3599 | 22.1821 | 29.5954 | 4.773% | 9.545% | 14.32% | 24.5943 | 27.1832
Pr, ppm | 1.2274 | 0.08268 | 1.062 | 1.3927 | 0.97932 | 1.4754 | 6.736% | 13.47% | 20.21% | 1.166 | 1.2887
Rb, ppm | 6.3303 | 0.44127 | 5.4477 | 7.2128 | 5.0064 | 7.6541 | 6.971% | 13.94% | 20.91% | 6.0137 | 6.6468
S, wt.% | 0.40163 | 0.015201 | 0.37122 | 0.43203 | 0.35602 | 0.44723 | 3.785% | 7.57% | 11.35% | 0.38154 | 0.42171
Sb, ppm | 0.8721 | 0.22564 | 0.42082 | 1.3234 | 0.19519 | 1.549 | 25.87% | 51.75% | 77.62% | 0.8285 | 0.91571
Sc, ppm | 4.9578 | 0.67654 | 3.6047 | 6.3108 | 2.9281 | 6.9874 | 13.65% | 27.29% | 40.94% | 4.7099 | 5.2056
Sm, ppm | 1.5174 | 0.1169 | 1.2836 | 1.7512 | 1.1667 | 1.8681 | 7.704% | 15.41% | 23.11% | 1.4415 | 1.5932
Sn, ppm | 0.65094 | 0.067766 | 0.5154 | 0.78647 | 0.44764 | 0.85424 | 10.41% | 20.82% | 31.23% | 0.61839 | 0.68348
Sr, ppm | 36.9431 | 6.5817 | 23.7797 | 50.1065 | 17.1979 | 56.6883 | 17.82% | 35.63% | 53.45% | 35.096 | 38.7903
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.30411 | 0.010198 | 0.28372 | 0.32451 | 0.27352 | 0.33471 | 3.353% | 6.707% | 10.06% | 0.28891 | 0.31932
Te, ppm | 0.10124 | 0.015981 | 0.069282 | 0.1332 | 0.053301 | 0.14919 | 15.78% | 31.57% | 47.35% | 0.096181 | 0.10631
Th, ppm | 0.80163 | 0.037641 | 0.72634 | 0.87691 | 0.6887 | 0.91455 | 4.696% | 9.391% | 14.09% | 0.76155 | 0.84171
Ti, wt.% | 0.31453 | 0.064105 | 0.18632 | 0.44274 | 0.12222 | 0.50685 | 20.38% | 40.76% | 61.14% | 0.29881 | 0.33026
Tl, ppm | 0.11831 | 0.0071241 | 0.10406 | 0.13255 | 0.096934 | 0.13968 | 6.022% | 12.04% | 18.07% | 0.11239 | 0.12422
Tm, ppm | 0.15719 | 0.026974 | 0.10324 | 0.21114 | 0.076265 | 0.23811 | 17.16% | 34.32% | 51.48% | 0.14933 | 0.16505
U, ppm | 0.20321 | 0.015604 | 0.172 | 0.23442 | 0.15639 | 0.25002 | 7.679% | 15.36% | 23.04% | 0.19305 | 0.21337
V, ppm | 127.4601 | 17.7831 | 91.894 | 163.0263 | 74.1109 | 180.8093 | 13.95% | 27.9% | 41.86% | 121.0871 | 133.8331
W, ppm | 23.5262 | 4.0047 | 15.5167 | 31.5356 | 11.512 | 35.5404 | 17.02% | 34.04% | 51.07% | 22.3499 | 24.7025
Y, ppm | 10.6429 | 0.53561 | 9.5717 | 11.7142 | 9.0361 | 12.2498 | 5.033% | 10.07% | 15.1% | 10.1108 | 11.1751
Yb, ppm | 0.97326 | 0.14022 | 0.69282 | 1.2537 | 0.5526 | 1.3939 | 14.41% | 28.81% | 43.22% | 0.92459 | 1.0219
Zn, ppm | 129.4807 | 6.8875 | 115.7058 | 143.2557 | 108.8184 | 150.1431 | 5.319% | 10.64% | 15.96% | 123.0067 | 135.9548
Zr, ppm | 13.9286 | 2.279 | 9.3705 | 18.4867 | 7.0914 | 20.7657 | 16.36% | 32.72% | 49.09% | 13.2322 | 14.625
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 234b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
N.A. | Not Applicable (Lab 1.27)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 234b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
Nagrom, Perth, WA, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 234b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 14 | 14 | 238557 | 0.08367 | 1.362 | 1.2975
2 | 1 | ANSLu | 1 | 9 | 9 | 238558 | 0.084944 | 1.3968 | 1.2994
2 | 1 | ANSLu | 1 | 19 | 19 | 238559 | 0.082549 | 1.2227 | 1.2901
2 | 1 | ANSLu | 1 | 13 | 13 | 238560 | 0.086222 | 1.31 | 1.2948
2 | 1 | ANSLu | 1 | 2 | 2 | 238561 | 0.082557 | 1.2849 | 1.2934
2 | 1 | ANSLu | 1 | 20 | 20 | 238562 | 0.085949 | 1.212 | 1.2896
2 | 1 | ANSLu | 1 | 7 | 7 | 238563 | 0.087758 | 1.3684 | 1.2979
2 | 1 | ANSLu | 1 | 18 | 18 | 238564 | 0.08677 | 1.296 | 1.294
2 | 1 | ANSLu | 1 | 10 | 10 | 238565 | 0.085683 | 1.2348 | 1.2908
2 | 1 | ANSLu | 1 | 12 | 12 | 238566 | 0.083603 | 1.3855 | 1.2988
2 | 1 | ANSLu | 1 | 3 | 3 | 238567 | 0.086673 | 1.2637 | 1.2923
2 | 1 | ANSLu | 1 | 6 | 6 | 238568 | 0.085909 | 1.2286 | 1.2904
2 | 1 | ANSLu | 1 | 17 | 17 | 238569 | 0.086368 | 1.2752 | 1.2929
2 | 1 | ANSLu | 1 | 15 | 15 | 238570 | 0.082662 | 1.2989 | 1.2942
2 | 1 | ANSLu | 1 | 5 | 5 | 238571 | 0.084608 | 1.3303 | 1.2959
2 | 1 | ANSLu | 1 | 4 | 4 | 238572 | 0.086652 | 1.2822 | 1.2933
2 | 1 | ANSLu | 1 | 11 | 11 | 238573 | 0.08751 | 1.2688 | 1.2926
2 | 1 | ANSLu | 1 | 16 | 16 | 238574 | 0.087577 | 1.2478 | 1.2915
2 | 1 | ANSLu | 1 | 1 | 1 | 238575 | 0.083323 | 1.2517 | 1.2917
2 | 1 | ANSLu | 1 | 8 | 8 | 238576 | 0.085381 | 1.3581 | 1.2973
Mean | [6 empty] | 0.085318 | 1.2939 | 1.2939
Median | [7 empty] | 1.2836 | 1.2934
SD | [7 empty] | 0.056397 | 0.0030076
RSD | [7 empty] | 4.359% | 0.2324%
RSD Upscale Factor | [7 empty] | 0.053329
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 234b (Certified Value 1.23 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 30g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 50g | 50g | 25g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 15g | 30g | 30g | 25g | 30g | 50g | 30g | 30g | 30g | 3
| 1 | 1 | 1.2517 | 1.022 | 1.259 | 1.386 | 1.2063 | 1.3053 | 1.2 | 1.26 | 1.28 | 1.14 | 1.28 | 1.222 | 1.322 | 1.161 | 1.26 | 1.26 | 1.245 | 1.196 | 1.185 | 1.19 | 1.18 | 1.22 | 1.2 | 1.189 | 1.2621 | 1.203 | 1.2928 | 1.267 | 1.22 | 1.244 | 1.235 | 1.37 | 1
| 1 | 2 | 1.2849 | 1.247 | 1.26 | 1.21 | 1.2273 | 1.278 | 1.24 | 1.24 | 1.22 | 1.21 | 1.28 | 1.232 | 1.3 | 1.155 | 1.255 | 1.25 | 1.25 | 1.242 | 1.2 | 1.192 | 1.23 | 1.231 | 1.2 | 1.197 | 1.2579 | 1.22 | 1.3088 | 1.265 | 1.21 | 1.236 | 1.201 | 1.34 | #N/A
| 1 | 3 | 1.2637 | 1.073 | 1.295 | 1.234 | 1.2211 | 1.2838 | 1.2 | 1.23 | 1.26 | 1.17 | 1.26 | 1.216 | 1.32 | 1.14 | 1.265 | 1.235 | 1.245 | 1.21 | 1.155 | 1.189 | 1.23 | 1.227 | 1.19 | 1.236 | 1.2558 | 1.21 | 1.2689 | 1.253 | 1.19 | 1.226 | 1.209 | 1.28 | 16
| 1 | 4 | 1.2822 | 1.071 | 1.235 | 1.234 | 1.2298 | 1.3053 | 1.22 | 1.23 | 1.24 | 1.17 | 1.28 | 1.233 | 1.336 | 1.052 | 1.25 | 1.265 | 1.235 | 1.284 | 1.23 | 1.184 | 1.17 | 1.232 | 1.18 | 1.207 | 1.2663 | 1.255 | 1.266 | 1.257 | 1.2 | 1.227 | 1.255 | 1.32 | 1.2342
| 1 | 5 | 1.3303 | 1.154 | 1.238 | 1.222 | 1.2264 | 1.2975 | 1.24 | 1.21 | 1.32 | 1.14 | 1.3 | 1.223 | 1.32 | 1.075 | 1.26 | 1.26 | 1.23 | 1.298 | 1.205 | 1.231 | 1.22 | 1.216 | 1.15 | 1.232 | 1.2537 | 1.205 | 1.3266 | 1.266 | 1.17 | 1.239 | 1.177 | 1.32 | 7
| 1 | 6 | 1.2286 | 1.126 | 1.239 | 1.238 | 1.2251 | 1.3012 | 1.2 | 1.26 | 1.2 | 1.25 | 1.27 | 1.203 | 1.336 | 1.154 | 1.25 | 1.28 | 1.235 | 1.294 | 1.215 | 1.213 | 1.19 | 1.27 | 1.19 | 1.263 | 1.2348 | 1.253 | 1.327 | 1.272 | 1.2 | 1.229 | 1.208 | 1.33
|  | 7 | 1.3684
|  | 8 | 1.3581
|  | 9 | 1.3968
|  | 10 | 1.2348
|  | 11 | 1.2688
|  | 12 | 1.3855
|  | 13 | 1.31
|  | 14 | 1.362
|  | 15 | 1.2989
|  | 16 | 1.2478
|  | 17 | 1.2752
|  | 18 | 1.296
|  | 19 | 1.2227
|  | 20 | 1.212
| Mean |  | 1.2939 | 1.1155 | 1.2543 | 1.254 | 1.2227 | 1.2952 | 1.2167 | 1.2383 | 1.2533 | 1.18 | 1.2783 | 1.2215 | 1.3223 | 1.1228 | 1.2567 | 1.2583 | 1.24 | 1.254 | 1.1983 | 1.1998 | 1.2033 | 1.2327 | 1.185 | 1.2207 | 1.2551 | 1.2243 | 1.2984 | 1.2633 | 1.1983 | 1.2335 | 1.2142 | 1.3267
| Median |  | 1.2836 | 1.0995 | 1.249 | 1.234 | 1.2257 | 1.2993 | 1.21 | 1.235 | 1.25 | 1.17 | 1.28 | 1.2225 | 1.321 | 1.147 | 1.2575 | 1.26 | 1.24 | 1.263 | 1.2025 | 1.191 | 1.205 | 1.229 | 1.19 | 1.2195 | 1.2569 | 1.215 | 1.3008 | 1.2655 | 1.2 | 1.2325 | 1.2085 | 1.325
| Std Dev. |  | 0.056397 | 0.079263 | 0.02273 | 0.065483 | 0.0085076 | 0.011573 | 0.019664 | 0.019408 | 0.043205 | 0.042895 | 0.013292 | 0.011113 | 0.013292 | 0.047038 | 0.0060553 | 0.015055 | 0.007746 | 0.044452 | 0.026013 | 0.01828 | 0.026583 | 0.019325 | 0.018708 | 0.027948 | 0.010925 | 0.023729 | 0.027097 | 0.0070048 | 0.017224 | 0.007287 | 0.027279 | 0.029439
| Rel.Std.Dev. |  | 4.359% | 7.106% | 1.812% | 5.222% | 0.6958% | 0.8935% | 1.616% | 1.567% | 3.447% | 3.635% | 1.04% | 0.9098% | 1.005% | 4.189% | 0.4819% | 1.196% | 0.6247% | 3.545% | 2.171% | 1.524% | 2.209% | 1.568% | 1.579% | 2.29% | 0.8705% | 1.938% | 2.087% | 0.5545% | 1.437% | 0.5908% | 2.247% | 2.219%
| PDM3 |  | 4.837% | -9.619% | 1.63% | 1.603% | -0.9355% | 4.939% | -1.422% | 0.3333% | 1.549% | -4.393% | 3.574% | -1.031% | 7.139% | -9.025% | 1.819% | 1.954% | 0.4684% | 1.603% | -2.908% | -2.786% | -2.502% | -0.1258% | -3.988% | -1.098% | 1.692% | -0.801% | 5.197% | 2.359% | -2.908% | -0.05827% | -1.625% | 7.49%
| Z-Score (Absolute) |  | NA | 3.43 | 0.61 | 0.6 | 0.32 | 1.8 | 0.49 | 0.14 | 0.58 | 1.56 | 1.3 | 0.35 | 2.59 | 3.22 | 0.67 | 0.72 | 0.19 | 0.6 | 1.02 | 0.98 | 0.88 | 0.02 | 1.41 | 0.37 | 0.63 | 0.27 | 1.89 | 0.87 | 1.02 | 0 | 0.56 | 2.71
| Analytical results for Pd in OREAS 234b (Indicative Value 18.6 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 1
|  |  | 25g | [31 empty] | 1
| 1 | 1 | 17.9185 | [31 empty] | 1
| 1 | 2 | 18.603 | [31 empty] | 1
| 1 | 3 | 19.9623 | [31 empty] | 16
| 1 | 4 | 17.884 | [31 empty] | 18.5625
| 1 | 5 | 17.6931 | [31 empty] | 7
| 1 | 6 | 19.3142
| Mean |  | 18.5625
| Median |  | 18.2608
| Std Dev. |  | 0.91183
| Rel.Std.Dev. |  | 4.912%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 234b (Indicative Value 10 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 1
|  |  | 25g | [31 empty] | 1
| 1 | 1 | 10.1638 | [31 empty] | 1
| 1 | 2 | 10.1075 | [31 empty] | 1
| 1 | 3 | 9.5345 | [31 empty] | 16
| 1 | 4 | 9.9442 | [31 empty] | 10.0121
| 1 | 5 | 10.0734 | [31 empty] | 7
| 1 | 6 | 10.2489
| Mean |  | 10.0121
| Median |  | 10.0905
| Std Dev. |  | 0.25481
| Rel.Std.Dev. |  | 2.545%
| PDM3 |  | -3.22e-13%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 234b (Certified Value 1.22 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | 5.05 | 5.06 | 5.07 | 6.01 | 6.02 | 6.03 | 6.04 | ppm
|  |  | INAA | Raw*PA | Various [a] | Raw*PA | Various [b] | Various [c] | Raw*PA | Various [d] | Raw*PA | Raw*PA | Raw*PA | Various [e] | 2
|  |  | 0.085g | 400g | see footer | 400g | see footer | see footer | 400g | see footer | 450g | 400g | 450g | see footer | 3
| 1 | 1 | 1.2517 | 1.2 | 1.31 |  | 1.291 | 1.197 | 1.1729 | 1.267 | 1.293 | 1.201 | 1.267 | 1.214 | 1
| 1 | 2 | 1.2849 | 1.17 | 1.34 |  | 1.2464 | 1.238 | 1.2773 | 1.209 | 1.149 | 1.254 | 1.223 | 1.208 | #N/A
| 1 | 3 | 1.2637 | 1.28 | 1.22 |  | 1.2142 | 1.155 | 1.1952 | 1.229 | 1.273 | 1.205 | 1.199 | 1.241 | 16
| 1 | 4 | 1.2822 | 1.17 | 1.32 |  | 1.2334 | 1.177 | 1.2095 | 1.279 | 1.193 | 1.262 | 1.255 | 1.252 | 1.2195
| 1 | 5 | 1.3303 | 1.19 | 1.24 |  | 1.2186 | 1.197 | 1.1696 | 1.282 | 1.23 | 1.291 | 1.161 | 1.261 | 9
| 1 | 6 | 1.2286 | 1.24 | 1.2 |  | 1.2585 | 1.243 | 1.1933 | 1.238 | 1.165 | 1.198 | 1.22 | 1.315
| 2 | 7 | 1.3684 | 1.24 | 1.21 |  | 1.17 | 1.19 | 1.207 | 1.134 | 1.228 | 1.132 | 1.224 | 1.08
| 2 | 8 | 1.3581 | 1.2 | 1.33 |  | 1.16 | 1.22 | 1.221 | 1.205 | 1.286 | 1.221 | 1.2416 | 1.201
| 2 | 9 | 1.3968 | 1.19 | 1.29 |  | 1.32 | 1.27 | 1.127 | 1.172 | 1.251 | 1.182 | 1.2011 | 1.202
| 2 | 10 | 1.2348 | 1.2 | 1.17 |  | 1.22 | 1.21 | 1.203 | 1.2 | 1.235 | 1.221 | 1.2542 | 1.271
| 2 | 11 | 1.2688 | 1.19 | 1.24 |  | 1.16 | 1.16 | 1.218 | 1.087 | 1.29 | 1.232 | 1.2005 | 1.214
| 2 | 12 | 1.3855 | 1.3 | 1.29 |  | 1.2 | 1.29 | 1.232 | 1.233 | 1.178 | 1.237 | 1.2386 | 1.096
| 3 | 13 | 1.31 | 1.32 | 1.28
| 3 | 14 | 1.362 | 1.25 | 1.24
| 3 | 15 | 1.2989 | 1.24 | 1.17
| 3 | 16 | 1.2478 | 1.32 | 1.31
| 3 | 17 | 1.2752 | 1.28 | 1.26
| 3 | 18 | 1.296 | 1.3 | 1.25
| 4 | 19 | 1.2227 | 1.17 | 1.18 | 1.26
| 4 | 20 | 1.212 | 1.25 | 1.3 | 1.13
| 4 | 21 |  | 1.22 | 1.22 | 1.24
| 4 | 22 |  | 1.23 | 1.28 | 1.16
| 4 | 23 |  | 1.22 | 1.17
| 4 | 24 |  | 1.27 | 1.2
| 5 | 25 |  |  | 1.14 | 1.22
| 5 | 26 |  |  | 1.15 | 1.17
| 5 | 27 |  |  | 1.26 | 1.25
| 5 | 28 |  |  | 1.19 | 1.15
| 5 | 29 |  |  | 1.17
| 5 | 30 |  |  | 1.19
| 6 | 31 |  |  | 1.19 | 1.22
| 6 | 32 |  |  | 1.24 | 1.25
| 6 | 33 |  |  | 1.22 | 1.13
| 6 | 34 |  |  | 1.29 | 1.25
| 6 | 35 |  |  | 1.29
| 6 | 36 |  |  | 1.27
| Mean |  | 1.2939 | 1.235 | 1.2394 | 1.2025 | 1.2243 | 1.2123 | 1.2021 | 1.2112 | 1.2309 | 1.2197 | 1.2237 | 1.2129
| Median |  | 1.2836 | 1.235 | 1.24 | 1.22 | 1.2193 | 1.2035 | 1.205 | 1.219 | 1.2325 | 1.221 | 1.2235 | 1.214
| Std Dev. |  | 0.056397 | 0.047365 | 0.055289 | 0.050655 | 0.049821 | 0.041848 | 0.036814 | 0.05825 | 0.050338 | 0.041179 | 0.030047 | 0.067452
| Rel.Std.Dev. |  | 4.359% | 3.835% | 4.461% | 4.212% | 4.069% | 3.452% | 3.062% | 4.809% | 4.089% | 3.376% | 2.455% | 5.561%
| PDM3 |  | 6.105% | 1.273% | 1.638% | -1.392% | 0.3995% | -0.5922% | -1.42% | -0.6742% | 0.9386% | 0.01603% | 0.3506% | -0.5375%
| Z-Score (Absolute) |  | NA | 1.23 | 1.58 | 1.38 | 0.37 | 0.59 | 1.4 | 0.67 | 0.9 | 0 | 0.33 | 0.54
[a]: Lab 5.02: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (350g), Rnd 03 - Raw*PA (350g), Rnd 04 - Raw*PA (400g), Rnd 05 - Raw*PA (400g), Rnd 06 - Raw*PA (400g)
[b]: Lab 5.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[c]: Lab 5.05: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (300g)
[d]: Lab 5.07: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[e]: Lab 6.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 234b (Certified Value 1.17 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.33 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.085g | 30g | 25g | 25g | 30g | 50g | 30g | 15g | 50g | 30g | 10g | 30g | 25g | 25g | 25g | 25g | 15g | 25g | 15g | 50g | 25g | 25g | 10g | 20g | 3
| 1 | 1 | 1.2517 | 1.16 | 1.32 | 1.2 | 1.232 | 1.2 | 1.11 | 1.19 | 1.1368 | 1.22 | 1.137 | 1.212 | 1.22 | 1.14 | 1.15 | 1.15 | 1.1817 | 1.189 | 1.0699 | 1.1791 | 1.225 | 1.05 | 1.235 | 1.1611 | 1
| 1 | 2 | 1.2849 | 1.13 | 1.32 | 1.2 | 1.214 | 1.22 | 1.1 | 1.31 | 1.144 | 1.25 | 1.165 | 1.252 | 1.22 | 1.16 | 1.15 | 1.16 | 1.0908 | 1.172 | 1.171 | 1.1272 | 1.16 | 1.04 | 1.236 | 1.1802 | #N/A
| 1 | 3 | 1.2637 | 1.16 | 1.32 | 1.12 | 1.205 | 1.19 | 1.1 | 1.29 | 1.1832 | 1.25 | 1.149 | 1.222 | 1.15 | 1.09 | 1.15 | 1.18 | 1.1716 | 1.208 | 1.1358 | 1.122 | 1.15 | 1.02 | 1.225 | 1.1614 | 16
| 1 | 4 | 1.2822 | 1.15 | 1.31 | 1.12 | 1.227 | 1.16 | 1.09 | 1.29 | 1.1616 | 1.27 | 1.189 | 1.226 | 1.2 | 1.14 | 1.18 | 1.18 | 1.1312 | 1.177 | 1.1452 | 1.1979 | 1.205 | 1.08 | 1.224 | 1.1747 | 1.1724
| 1 | 5 | 1.3303 | 1.14 | 1.32 | 1.12 | 1.231 | 1.24 | 1.07 | 1.3 | 1.1648 | 1.26 | 1.147 | 1.252 | 1.18 | 1.13 | 1.2 | 1.18 | 1.1211 | 1.22 | 1.1387 | 1.1087 | 1.155 | 1.04 | 1.231 | 1.1637 | 11
| 1 | 6 | 1.2286 | 1.15 | 1.29 | 1.12 | 1.213 | 1.16 | 1.2 | 1.29 | 1.17 | 1.28 | 1.131 | 1.2328 | 1.12 | 1.13 | 1.17 | 1.14 | 1.1413 | 1.209 | 1.0067 | 1.1265 | 1.14 | 1.08 | 1.244 | 1.1484
|  | 7 | 1.3684
|  | 8 | 1.3581
|  | 9 | 1.3968
|  | 10 | 1.2348
|  | 11 | 1.2688
|  | 12 | 1.3855
|  | 13 | 1.31
|  | 14 | 1.362
|  | 15 | 1.2989
|  | 16 | 1.2478
|  | 17 | 1.2752
|  | 18 | 1.296
|  | 19 | 1.2227
|  | 20 | 1.212
| Mean |  | 1.2939 | 1.1483 | 1.3133 | 1.1467 | 1.2203 | 1.195 | 1.1117 | 1.2783 | 1.1601 | 1.255 | 1.153 | 1.2328 | 1.1817 | 1.1317 | 1.1667 | 1.165 | 1.1396 | 1.1958 | 1.1112 | 1.1436 | 1.1725 | 1.0517 | 1.2325 | 1.1649
| Median |  | 1.2836 | 1.15 | 1.32 | 1.12 | 1.2205 | 1.195 | 1.1 | 1.29 | 1.1632 | 1.255 | 1.148 | 1.2294 | 1.19 | 1.135 | 1.16 | 1.17 | 1.1362 | 1.1985 | 1.1372 | 1.1269 | 1.1575 | 1.045 | 1.233 | 1.1625
| Std Dev. |  | 0.056397 | 0.01169 | 0.012111 | 0.041312 | 0.011165 | 0.032094 | 0.04535 | 0.044008 | 0.017077 | 0.020736 | 0.021128 | 0.016327 | 0.040208 | 0.023166 | 0.020656 | 0.017607 | 0.033447 | 0.019364 | 0.061187 | 0.03593 | 0.034169 | 0.024014 | 0.0075033 | 0.011229
| Rel.Std.Dev. |  | 4.359% | 1.018% | 0.9221% | 3.603% | 0.9149% | 2.686% | 4.08% | 3.443% | 1.472% | 1.652% | 1.832% | 1.324% | 3.403% | 2.047% | 1.771% | 1.511% | 2.935% | 1.619% | 5.506% | 3.142% | 2.914% | 2.283% | 0.6088% | 0.964%
| PDM3 |  | 10.36% | -2.057% | 12.02% | -2.199% | 4.084% | 1.924% | -5.184% | 9.031% | -1.056% | 7.041% | -1.659% | 5.148% | 0.7864% | -3.478% | -0.4929% | -0.6351% | -2.8% | 1.995% | -5.222% | -2.463% | 0.004606% | -10.3% | 5.122% | -0.6422%
| Z-Score (Absolute) |  | NA | 0.37 | 3.33 | 0.41 | 1.24 | 0.67 | 1.2 | 2.55 | 0.11 | 2.02 | 0.27 | 1.52 | 0.37 | 0.75 | 0.04 | 0 | 0.57 | 0.69 | 1.21 | 0.48 | 0.17 | 2.55 | 1.52 | 0
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 234b (Certified Value 1.17 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.30 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 05g | 50g | 60g | 30g | 30g | 30g | 200g | 30g | 200g | 20g | 30g | 05g | 3
| 1 | 1 | 1.2517 | 1.037 | 1.18 | 1.075 | 1.096 | 1.0491 | 1.1 | 1.34 | 1.22 | 1.2 | 1.166 | 1.15 | 1.28 | 1.09 | 1.11 | 1.23 | 1.25 | 1.1857 | 1.19 | 1.215 | 1.18 | 1
| 1 | 2 | 1.2849 | 1.063 | 1.18 | 1.1 | 1.088 | 1.0886 | 1.14 | 1.23 | 1.18 | 1.2 | 1.16 | 1.1 | 1.31 | 0.72 | 1.21 | 1.19 | 1.23 | 1.1769 | 1.18 | 1.216 | 1.2 | #N/A
| 1 | 3 | 1.2637 | 1.059 | 1.2 | 1.15 | 1.128 | 1.0789 | 1.08 | 1.19 | 1.14 | 1.2 | 1.1799 | 1.14 | 1.29 | 1.1 | 1.15 | 1.19 | 1.24 | 1.1748 | 1.17 | 1.216 | 1.18 | 16
| 1 | 4 | 1.2822 | 1.039 | 1.25 | 1.1 | 1.12 | 1.0631 | 1.12 | 1.17 | 1.19 | 1.2 | 1.1659 | 1.12 | 1.28 | 0.7 | 1.17 | 1.23 | 1.23 | 1.183 | 1.16 | 1.208 | 1.17 | 1.1731
| 1 | 5 | 1.3303 | 1.018 | 1.21 | 1.125 | 1.048 | 1.0946 | 1.13 | 1.22 | 1.23 | 1.2 | 1.1619 | 1.14 | 1.32 | 1.2 | 1.13 | 1.22 | 1.15 | 1.1657 | 1.18 | 1.207 | 1.18 | 13
| 1 | 6 | 1.2286 | 1.061 | 1.21 | 1.125 | 1.072 | 1.0842 | 1.08 | 1.23 | 1.18 | 1.1 | 1.19 | 1.09 | 1.3 | 1.11 | 1.29 | 1.213 | 1.17 | 1.1751 | 1.18 | 1.198 | 1.2
|  | 7 | 1.3684
|  | 8 | 1.3581
|  | 9 | 1.3968
|  | 10 | 1.2348
|  | 11 | 1.2688
|  | 12 | 1.3855
|  | 13 | 1.31
|  | 14 | 1.362
|  | 15 | 1.2989
|  | 16 | 1.2478
|  | 17 | 1.2752
|  | 18 | 1.296
|  | 19 | 1.2227
|  | 20 | 1.212
| Mean |  | 1.2939 | 1.0462 | 1.205 | 1.1125 | 1.092 | 1.0764 | 1.1083 | 1.23 | 1.19 | 1.1833 | 1.1706 | 1.1233 | 1.2967 | 0.98667 | 1.1767 | 1.2122 | 1.2117 | 1.1769 | 1.1767 | 1.21 | 1.185
| Median |  | 1.2836 | 1.049 | 1.205 | 1.1125 | 1.092 | 1.0816 | 1.11 | 1.225 | 1.185 | 1.2 | 1.166 | 1.13 | 1.295 | 1.095 | 1.16 | 1.2165 | 1.23 | 1.176 | 1.18 | 1.2115 | 1.18
| Std Dev. |  | 0.056397 | 0.017871 | 0.025884 | 0.02622 | 0.029826 | 0.017129 | 0.025626 | 0.058992 | 0.032249 | 0.040825 | 0.011781 | 0.024221 | 0.01633 | 0.21796 | 0.06532 | 0.018335 | 0.041191 | 0.007029 | 0.010328 | 0.0071274 | 0.012247
| Rel.Std.Dev. |  | 4.359% | 1.708% | 2.148% | 2.357% | 2.731% | 1.591% | 2.312% | 4.796% | 2.71% | 3.45% | 1.006% | 2.156% | 1.259% | 22.09% | 5.551% | 1.513% | 3.4% | 0.5973% | 0.8777% | 0.589% | 1.034%
| PDM3 |  | 10.3% | -10.82% | 2.722% | -5.163% | -6.911% | -8.239% | -5.519% | 4.853% | 1.443% | 0.875% | -0.2081% | -4.24% | 10.54% | -15.89% | 0.3067% | 3.333% | 3.29% | 0.3236% | 0.3067% | 3.148% | 1.017%
| Z-Score (Absolute) |  | NA | 2.51 | 0.54 | 1.23 | 1.63 | 1.92 | 1.31 | 1.02 | 0.25 | 0.13 | 0.12 | 1.03 | 2.3 | 3.65 | 0 | 0.68 | 0.67 | 0 | 0 | 0.64 | 0.16
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 234b (Certified Value 0.348 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.5 | 0.34 | 0.33142 | 0.4 | 0.46 | 0.4 | 0.31 | < 0.5 | 0.3 | 0.36 | 0.38 | 0.35 | 0.36 | 0.2 | 0.3 | 0.3 | 0.35 | 0.21 | < 1 | 0.35 | 0.3306 | 0.33 | 0.36 | 0.4 | 0.3 |  |  |  |  | 1
| 1 | 2 | 0.52 | 0.37 | 0.34823 | 0.4 | 0.42 | 0.4 | 0.36 | < 0.5 | 0.4 | 0.34 | 0.36 | 0.36 | 0.33 | 0.3 | 0.3 | 0.3 | 0.34 | 0.23 | < 1 | 0.34 | 0.3213 | 0.33 | 0.35 | 0.39 | 0.31 |  |  |  |  | 19
| 1 | 3 | 0.49 | 0.36 | 0.34846 | 0.3 | 0.39 | 0.3 | 0.33 | < 0.5 | 0.4 | 0.37 | 0.36 | 0.34 | 0.34 | 0.2 | 0.3 | 0.3 | 0.35 | 0.36 | < 1 | 0.35 | 0.34 | 0.33 | 0.34 | 0.37 | 0.32 |  |  |  |  | 16
| 1 | 4 | 0.48 | 0.37 | 0.33854 | 0.4 | 0.43 | 0.4 | 0.36 | < 0.5 | 0.3 | 0.33 | 0.36 | 0.34 | 0.36 | 0.3 | 0.3 | 0.3 | 0.38 | 0.28 | < 1 | 0.35 | 0.3335 | 0.36 | 0.32 | 0.4 | 0.33 |  |  |  |  | 0.34803
| 1 | 5 | 0.51 | 0.34 | 0.32263 | 0.4 | 0.41 | 0.4 | 0.33 | < 0.5 | 0.3 | 0.35 | 0.38 | 0.36 | 0.34 | 0.2 | 0.3 | 0.3 | 0.34 | 0.29 | < 1 | 0.34 | 0.3203 | 0.36 | 0.33 | 0.38 | 0.32 |  |  |  |  | 15
| 1 | 6 | 0.48 | 0.34 | 0.34346 | < 0.3 | 0.41 | 0.4 | 0.33 | < 0.5 | 0.4 | 0.35 | 0.37 | 0.36 | 0.37 | < 0.2 | 0.3 | 0.3 | 0.37 | 0.29 | < 1 | 0.36 | 0.3158 | 0.33 | 0.37 | 0.4 | 0.33
| Mean |  | 0.49667 | 0.35333 | 0.33879 | 0.38 | 0.42 | 0.38333 | 0.33667 |  | 0.35 | 0.35 | 0.36833 | 0.35167 | 0.35 | 0.24 | 0.3 | 0.3 | 0.355 | 0.27667 |  | 0.34833 | 0.32692 | 0.34 | 0.345 | 0.39 | 0.31833
| Median |  | 0.495 | 0.35 | 0.341 | 0.4 | 0.415 | 0.4 | 0.33 |  | 0.35 | 0.35 | 0.365 | 0.355 | 0.35 | 0.2 | 0.3 | 0.3 | 0.35 | 0.285 |  | 0.35 | 0.32595 | 0.33 | 0.345 | 0.395 | 0.32
| Std Dev. |  | 0.01633 | 0.015055 | 0.010199 | 0.044721 | 0.023664 | 0.040825 | 0.019664 |  | 0.054772 | 0.014142 | 0.0098319 | 0.0098319 | 0.015492 | 0.054772 | 0 | 0 | 0.016432 | 0.052789 |  | 0.0075277 | 0.0092411 | 0.015492 | 0.018708 | 0.012649 | 0.01169
| Rel.Std.Dev. |  | 3.288% | 4.261% | 3.01% | 11.77% | 5.634% | 10.65% | 5.841% |  | 15.65% | 4.041% | 2.669% | 2.796% | 4.426% | 22.82% | 0% | 0% | 4.629% | 19.08% |  | 2.161% | 2.827% | 4.556% | 5.423% | 3.243% | 3.672%
| PDM3 |  | 42.71% | 1.525% | -2.654% | 9.187% | 20.68% | 10.14% | -3.264% |  | 0.567% | 0.567% | 5.835% | 1.046% | 0.567% | -31.04% | -13.8% | -13.8% | 2.004% | -20.5% |  | 0.0881% | -6.066% | -2.306% | -0.8697% | 12.06% | -8.532%
| Z-Score (Absolute) |  | 7.42 | 0.17 | 0.57 | NA | 3.54 | NA | 0.67 | 5.06 | NA | 0 | 0.93 | 0.08 | 0 | NA | NA | NA | 0.25 | 3.71 | 7.59 | 0.08 | 1.17 | 0.51 | 0.25 | 2.02 | 1.6
Results from laboratories 1.07, 1.10, 1.13, 1.19, 1.22 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 234b (Certified Value 6.72 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.68 | 6.5988 | 6.9479 | 6.4888 | 6.51 | 7.568 | 7.02 | 6.3309 | 6.36 | 6.47 | 7 | 6.97 | 6.94 | 6.52 | 7 | 6.729 | 6.76 | 6.81 | 5.19 | 6.7007 | 6.635 | 6.69 | 6.477 | 6.89 | 7.25 | 6.95 | 6.9 |  |  | 1
| 1 | 2 | 6.99 | 6.602 | 7.0008 | 6.795 | 6.6 | 7.621 | 6.88 | 6.6433 | 6.46 | 6.8 | 6.83 | 6.79 | 6.4 | 6.55 | 6.96 | 6.874 | 6.84 | 6.57 | 4.11 | 6.5024 | 6.656 | 6.67 | 6.469 | 6.67 | 6.98 | 6.93 | 6.7 |  |  | #N/A
| 1 | 3 | 7.3 | 6.5584 | 6.9868 | 6.9906 | 6.66 | 7.409 | 7.01 | 6.3804 | 6.4 | 6.76 | 6.96 | 6.6 | 6.64 | 6.6 | 6.87 | 6.805 | 6.8 | 6.51 | 4.25 | 6.325 | 6.628 | 6.71 | 6.517 | 6.69 | 7.05 | 6.99 | 6.9 |  |  | 16
| 1 | 4 | 7.23 | 6.5585 | 6.9478 | 6.7392 | 6.6 | 7.409 | 6.92 | 6.534 | 6.42 | 6.81 | 6.83 | 6.59 | 6.35 | 6.5 | 6.55 | 6.766 | 6.8 | 6.53 | 3.89 | 6.5429 | 6.591 | 6.68 | 6.3 | 6.71 | 7.11 | 6.95 | 6.9 |  |  | 6.7183
| 1 | 5 | 7.04 | 6.5988 | 6.9953 | 6.0971 | 6.63 | 7.515 | 7.02 | 6.527 | 6.47 | 6.78 | 6.74 | 6.65 | 6.83 | 6.64 | 6.85 | 6.921 | 6.59 | 6.91 | 4.48 | 6.5231 | 6.648 | 6.75 | 6.533 | 6.83 | 7.18 | 6.9 | 6.2 |  |  | 16
| 1 | 6 | 7.5 | 6.5988 | 6.9812 | 6.6065 | 6.6 | 7.356 | 6.84 | 6.41 | 6.37 | 6.65 | 6.85 | 6.76 | 6.62 | 6.61 | 6.34 | 6.651 | 6.73 | 6.67 | 4 | 6.3498 | 6.618 | 6.51 | 6.479 | 6.69 | 7.07 | 6.94 | 7.8
| Mean |  | 7.1233 | 6.5859 | 6.9766 | 6.6195 | 6.6 | 7.4797 | 6.9483 | 6.4709 | 6.4133 | 6.7117 | 6.8683 | 6.7267 | 6.63 | 6.57 | 6.7617 | 6.791 | 6.7533 | 6.6667 | 4.32 | 6.4906 | 6.6293 | 6.6683 | 6.4625 | 6.7467 | 7.1067 | 6.9433 | 6.9
| Median |  | 7.135 | 6.5988 | 6.984 | 6.6728 | 6.6 | 7.462 | 6.965 | 6.4685 | 6.41 | 6.77 | 6.84 | 6.705 | 6.63 | 6.575 | 6.86 | 6.7855 | 6.78 | 6.62 | 4.18 | 6.5127 | 6.6315 | 6.685 | 6.478 | 6.7 | 7.09 | 6.945 | 6.9
| Std Dev. |  | 0.28486 | 0.021286 | 0.023325 | 0.30756 | 0.0502 | 0.10422 | 0.079099 | 0.11692 | 0.045461 | 0.13167 | 0.095376 | 0.14487 | 0.23134 | 0.055136 | 0.26011 | 0.098035 | 0.088468 | 0.16269 | 0.47303 | 0.13807 | 0.023201 | 0.082563 | 0.083512 | 0.090701 | 0.096471 | 0.029439 | 0.51769
| Rel.Std.Dev. |  | 3.999% | 0.3232% | 0.3343% | 4.646% | 0.7606% | 1.393% | 1.138% | 1.807% | 0.7088% | 1.962% | 1.389% | 2.154% | 3.489% | 0.8392% | 3.847% | 1.444% | 1.31% | 2.44% | 10.95% | 2.127% | 0.35% | 1.238% | 1.292% | 1.344% | 1.357% | 0.424% | 7.503%
| PDM3 |  | 6.029% | -1.971% | 3.845% | -1.47% | -1.761% | 11.33% | 3.424% | -3.682% | -4.54% | -0.09891% | 2.233% | 0.1244% | -1.314% | -2.208% | 0.6453% | 1.082% | 0.5213% | -0.7687% | -35.7% | -3.389% | -1.324% | -0.7439% | -3.808% | 0.4221% | 5.781% | 3.349% | 2.704%
| Z-Score (Absolute) |  | 1.96 | 0.6 | 1.26 | 0.44 | 0.53 | 3.66 | 1.13 | 1.15 | 1.42 | 0 | 0.75 | 0.07 | 0.39 | 0.67 | 0.24 | 0.38 | 0.2 | 0.21 | 11.38 | 1.05 | 0.39 | 0.21 | 1.19 | 0.17 | 1.88 | 1.1 | 0.9
| Analytical results for As in OREAS 234b (Certified Value 65 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 77 | 66.8 | 66.8306 | 57.3 | 67 | 61.8 | 64 | 54.8 | 68 | 68.4 | 66.5 | 67.8 | 68 | 63 | 62 | 67 | 41.6 | 63 | 67.1 | 65.6279 | 67.5 | 70 | 66.7 | 58.2 | [5 empty] | 1
| 1 | 2 | 77 | 68.3 | 70.3632 | 64.8 | 66 | 66.1 | 60 | 55.8 | 65.4 | 70.5 | 65.6 | 66.8 | 67 | 63 | 63 | 64 | 49.91 | 65 | 66.5 | 64.1866 | 67.2 | 66 | 66 | 57.6 | [5 empty] | 20
| 1 | 3 | 76 | 68.6 | 70.8297 | 68.4 | 67 | 64.6 | 62 | 54.4 | 66.5 | 66.4 | 65.7 | 66.1 | 67 | 62 | 63 | 65 | 52.74 | 66 | 67.4 | 63.9374 | 68.5 | 65 | 64.4 | 59.7 | [5 empty] | 16
| 1 | 4 | 77 | 66.8 | 69.4982 | 63.4 | 67 | 66.6 | 60 | 50.8 | 64.9 | 67.4 | 65 | 65.4 | 64 | 59 | 63 | 65 | 51.029 | 64 | 67.9 | 61.1471 | 68.8 | 67 | 63.2 | 63.1 | [5 empty] | 65.2859
| 1 | 5 | 74 | 67.9 | 66.0658 | 62 | 68 | 65 | 62 | 51.1 | 63.6 | 66.9 | 69.4 | 66.3 | 67 | 60 | 62 | 68 | 49.83 | 66 | 67.5 | 61.9635 | 69.7 | 68 | 64.1 | 60.6 | [5 empty] | 17
| 1 | 6 | 79 | 64.7 | 69.6131 | 42.3 | 64 | 63.1 | 60 | 52.1 | 65.6 | 68.5 | 71.6 | 67.9 | 63 | 61 | 62 | 65 | 49.21 | 62 | 65.5 | 61.979 | 68.6 | 67 | 64.9 | 65.7
| Mean |  | 76.6667 | 67.1833 | 68.8668 | 59.7 | 66.5 | 64.5333 | 61.3333 | 53.1667 | 65.6667 | 68.0167 | 67.3 | 66.7167 | 66 | 61.3333 | 62.5 | 65.6667 | 49.0532 | 64.3333 | 66.9833 | 63.1403 | 68.3833 | 67.1667 | 64.8833 | 60.8167
| Median |  | 77 | 67.35 | 69.5556 | 62.7 | 67 | 64.8 | 61 | 53.25 | 65.5 | 67.9 | 66.1 | 66.55 | 67 | 61.5 | 62.5 | 65 | 49.87 | 64.5 | 67.25 | 62.9582 | 68.55 | 67 | 64.65 | 60.15
| Std Dev. |  | 1.633 | 1.4303 | 1.9514 | 9.265 | 1.3784 | 1.8151 | 1.633 | 2.104 | 1.4882 | 1.4689 | 2.6184 | 0.98675 | 2 | 1.633 | 0.54772 | 1.5055 | 3.8577 | 1.633 | 0.86352 | 1.7102 | 0.91086 | 1.7224 | 1.2828 | 3.0838
| Rel.Std.Dev. |  | 2.13% | 2.129% | 2.834% | 15.52% | 2.073% | 2.813% | 2.662% | 3.957% | 2.266% | 2.16% | 3.891% | 1.479% | 3.03% | 2.662% | 0.8764% | 2.293% | 7.864% | 2.538% | 1.289% | 2.709% | 1.332% | 2.564% | 1.977% | 5.071%
| PDM3 |  | 17.43% | 2.906% | 5.485% | -8.556% | 1.86% | -1.153% | -6.054% | -18.56% | 0.5832% | 4.183% | 3.085% | 2.192% | 1.094% | -6.054% | -4.267% | 0.5832% | -24.86% | -1.459% | 2.6% | -3.287% | 4.744% | 2.881% | -0.6166% | -6.846%
| Z-Score (Absolute) |  | 3.72 | 0.51 | 1.08 | 2.02 | 0.28 | 0.38 | 1.47 | 4.23 | 0 | 0.8 | 0.55 | 0.36 | 0.11 | 1.47 | 1.07 | 0 | 5.62 | 0.45 | 0.45 | 0.86 | 0.92 | 0.51 | 0.27 | 1.64
| Analytical results for B in OREAS 234b (Indicative Value 47.6 ppm) | [30 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.03 | 1.25 | 1.27 | [26 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 1
1
| 1 | 1 | 445.22 | 11280.37 | 10 | [26 empty] | 1
| 1 | 2 | 205.08 | 15629.91 | 10 | [26 empty] | 3
| 1 | 3 | 54.45 | 18803.15 | 10 | [26 empty] | 16
| 1 | 4 | 55.72 | 16767.63 | 5 | [26 empty] | 47.6333
| 1 | 5 | 58.17 | 17111.12 | 10 | [26 empty] | 9
| 1 | 6 | 57.08 | 16888.68 | 10
| Mean |  | 145.9533 | 16080.1433 | 9.1667
| Median |  | 57.625 | 16828.155 | 10
| Std Dev. |  | 158.225 | 2562.9897 | 2.0412
| Rel.Std.Dev. |  | 108.4% | 15.94% | 22.27%
| PDM3 |  | 206.4% | 33658% | -80.76%
| Z-Score (Absolute) |  | 0 | 78.55 | 0.67
| Analytical results for Ba in OREAS 234b (Certified Value 290 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 295 | 298.2 | 303.3453 | 270.568 | 270 | 336 | 290 | 284.9 | 270 | 284 | 310 | 300 | 300 | 290 | 239 | 286 | 273 | 276 | 810.11 | 284.04 | 290 | 288.334 | 290 | 319.9108 | 299 | 328 | 281 | 269 |  | 1
| 1 | 2 | 292 | 298.7 | 303.2114 | 294.4646 | 287 | 336 | 287 | 293.3 | 269 | 291 | 300 | 300 | 310 | 290 | 238 | 285 | 270 | 265 | 1195.48 | 273.474 | 290 | 290.0865 | 290 | 319.3447 | 288 | 309 | 279 | 265 |  | 21
| 1 | 3 | 294 | 289.6 | 303.7058 | 297.8355 | 284 | 328 | 292 | 287 | 270 | 295 | 310 | 280 | 290 | 290 | 239 | 287 | 268 | 264 | 1554.36 | 266.049 | 290 | 290.1382 | 300 | 323.7037 | 292 | 321 | 282 | 279 |  | 16
| 1 | 4 | 296 | 295.9 | 301.0278 | 293.2387 | 262 | 327 | 293 | 295.9 | 267 | 291 | 300 | 290 | 280 | 290 | 223 | 288 | 278 | 265 | 1503.91 | 276.129 | 285 | 287.831 | 290 | 310.6064 | 291 | 336 | 281 | 309 |  | 289.8811
| 1 | 5 | 285 | 290.3 | 305.4465 | 271.7073 | 268 | 331 | 293 | 292.6 | 268 | 301 | 300 | 290 | 310 | 290 | 235 | 287 | 266 | 279 | 1430.07 | 275.094 | 290 | 292.3663 | 300 | 314.6651 | 297 | 317 | 278 | 290 |  | 18
| 1 | 6 | 302 | 286.4 | 309.9167 | 289.2023 | 271 | 328 | 293 | 294.9 | 268 | 291 | 300 | 290 | 290 | 300 | 220 | 286 | 267 | 271 | 1213.79 | 266.85 | 285 | 292.5659 | 290 | 313.102 | 292 | 321 | 281 | 330
| Mean |  | 294 | 293.1833 | 304.4423 | 286.1694 | 273.6667 | 331 | 291.3333 | 291.4333 | 268.6667 | 292.1667 | 303.3333 | 291.6667 | 296.6667 | 291.6667 | 232.3333 | 286.5 | 270.3333 | 270 | 1284.62 | 273.606 | 288.3333 | 290.2203 | 293.3333 | 316.8888 | 293.1667 | 322 | 280.3333 | 290.3333
| Median |  | 294.5 | 293.1 | 303.5256 | 291.2205 | 270.5 | 329.5 | 292.5 | 292.95 | 268.5 | 291 | 300 | 290 | 295 | 290 | 236.5 | 286.5 | 269 | 268 | 1321.93 | 274.284 | 290 | 290.1123 | 290 | 317.0049 | 292 | 321 | 281 | 284.5
| Std Dev. |  | 5.5498 | 5.1019 | 3.0298 | 11.972 | 9.7297 | 4.0988 | 2.4221 | 4.4532 | 1.2111 | 5.6006 | 5.164 | 7.5277 | 12.1106 | 4.0825 | 8.5713 | 1.0488 | 4.5019 | 6.3875 | 275.7267 | 6.6365 | 2.582 | 1.9697 | 5.164 | 4.9063 | 4.0702 | 9.252 | 1.5055 | 25.0892
| Rel.Std.Dev. |  | 1.888% | 1.74% | 0.9952% | 4.184% | 3.555% | 1.238% | 0.8314% | 1.528% | 0.4508% | 1.917% | 1.702% | 2.581% | 4.082% | 1.4% | 3.689% | 0.3661% | 1.665% | 2.366% | 21.46% | 2.426% | 0.8955% | 0.6787% | 1.76% | 1.548% | 1.388% | 2.873% | 0.5371% | 8.642%
| PDM3 |  | 1.421% | 1.139% | 5.023% | -1.28% | -5.593% | 14.18% | 0.501% | 0.5355% | -7.318% | 0.7885% | 4.641% | 0.616% | 2.341% | 0.616% | -19.85% | -1.166% | -6.743% | -6.858% | 343.2% | -5.614% | -0.5339% | 0.117% | 1.191% | 9.317% | 1.133% | 11.08% | -3.294% | 0.156%
| Z-Score (Absolute) |  | 0.31 | 0.21 | 1.66 | 0.69 | 2.3 | 5.07 | 0.03 | 0.01 | 2.94 | 0.08 | 1.51 | 0.01 | 0.66 | 0.01 | 7.61 | 0.65 | 2.73 | 2.77 | 127.58 | 2.31 | 0.41 | 0.17 | 0.23 | 3.26 | 0.21 | 3.91 | 1.44 | 0.16
| Analytical results for Be in OREAS 234b (Certified Value 0.44 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.4 | 0.5 | 0.31 | 0.4045 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.42 | < 5 | < 1 | 0.46 | 0.46 | 0.47 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.43 | < 0.5 | 0.41928 | 0.4 | 0.33771 | < 0.05 | 0.5 | 0.45 | 0.38 |  |  | 1
| 1 | 2 | 0.4 | 0.51 | 0.31 | 0.38914 | < 0.5 | 0.4 | 0.5 | 0.5 | 0.52 | < 5 | < 1 | 0.43 | 0.46 | 0.45 | 0.41 | < 0.5 | < 0.5 | < 1 | 0.38 | 0.5 | 0.42453 | 0.4 | 0.32995 | < 0.05 | 0.5 | 0.48 | 0.35 |  |  | 22
| 1 | 3 | 0.5 | 0.45 | 0.3 | 0.39073 | < 0.5 | 0.5 | 0.5 | 0.6 | 0.41 | < 5 | < 1 | 0.44 | 0.43 | 0.45 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.4 | < 0.5 | 0.42134 | 0.4 | 0.33632 | 0.08 | 0.5 | 0.47 | 0.34 |  |  | 16
| 1 | 4 | 0.5 | 0.47 | 0.29 | 0.4025 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.48 | < 5 | < 1 | 0.44 | 0.42 | 0.44 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.42 | < 0.5 | 0.43858 | 0.39 | 0.32613 | < 0.05 | 0.5 | 0.46 | 0.58 |  |  | 0.43576
| 1 | 5 | 0.5 | 0.47 | 0.29 | 0.38773 | < 0.5 | 0.4 | 0.5 | 0.6 | 0.4 | < 5 | < 1 | 0.43 | 0.44 | 0.45 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.39 | < 0.5 | 0.42742 | 0.4 | 0.33637 | < 0.05 | 0.5 | 0.43 | 0.45 |  |  | 19
| 1 | 6 | 0.5 | 0.44 | 0.27 | 0.3983 | < 0.5 | 0.5 | 0.5 | 0.6 | 0.51 | < 5 | < 1 | 0.43 | 0.44 | 0.46 | 0.4 | < 0.5 | < 0.5 | < 1 | 0.41 | < 0.5 | 0.42375 | 0.38 | 0.33349 | < 0.05 | 0.5 | 0.45 | 0.33
| Mean |  | 0.46667 | 0.47333 | 0.295 | 0.39548 |  | 0.43333 | 0.5 | 0.58333 | 0.45667 |  |  | 0.43833 | 0.44167 | 0.45333 | 0.40167 |  |  |  | 0.405 | 0.5 | 0.42582 | 0.395 | 0.33333 | 0.08 | 0.5 | 0.45667 | 0.405
| Median |  | 0.5 | 0.47 | 0.295 | 0.39451 |  | 0.4 | 0.5 | 0.6 | 0.45 |  |  | 0.435 | 0.44 | 0.45 | 0.4 |  |  |  | 0.405 | 0.5 | 0.42414 | 0.4 | 0.3349 | 0.08 | 0.5 | 0.455 | 0.365
| Std Dev. |  | 0.05164 | 0.027325 | 0.015166 | 0.007232 |  | 0.05164 | 0 | 0.040825 | 0.053166 |  |  | 0.01169 | 0.016021 | 0.010328 | 0.0040825 |  |  |  | 0.018708 |  | 0.0068435 | 0.0083666 | 0.0044887 |  | 0 | 0.017512 | 0.096073
| Rel.Std.Dev. |  | 11.07% | 5.773% | 5.141% | 1.829% |  | 11.92% | 0% | 6.999% | 11.64% |  |  | 2.667% | 3.627% | 2.278% | 1.016% |  |  |  | 4.619% |  | 1.607% | 2.118% | 1.347% |  | 0% | 3.835% | 23.72%
| PDM3 |  | 7.092% | 8.622% | -32.3% | -9.244% |  | -0.5577% | 14.74% | 33.86% | 4.797% |  |  | 0.5898% | 1.355% | 4.032% | -7.825% |  |  |  | -7.06% | 14.74% | -2.283% | -9.354% | -23.51% | -81.64% | 14.74% | 4.797% | -7.06%
| Z-Score (Absolute) |  | 0.56 | 0.67 | 2.33 | 0.64 | 3.09 | 0 | 1.12 | 2.53 | 0.39 | 34.84 | 1.12 | 0.08 | 0.14 | 0.34 | 0.53 | 3.09 | 3.09 | 1.12 | 0.48 | 2.39 | 0.13 | 0.65 | 1.69 | 6.73 | 1.12 | 0.39 | 0.48
| Analytical results for Bi in OREAS 234b (Certified Value 0.047 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.05 | 0.04 | 0.069459 | 0.05 | 0.2 | 0.07 | 0.05 | < 0.1 | 0.04 | 0.05 | 0.05 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.05 | 0.0184 | 0.05 | 0.06 | 0.05 | 0.03 |  |  |  |  | 1
| 1 | 2 | 0.06 | 0.05 | 0.072081 | 0.05 | 0.2 | 0.07 | 0.05 | < 0.1 | 0.05 | 0.04 | 0.05 | 0.03 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.04 | 0.016 | 0.05 | 0.06 | 0.05 | 0.03 |  |  |  |  | 23
| 1 | 3 | 0.06 | 0.04 | 0.069148 | 0.05 | 0.2 | 0.07 | 0.05 | < 0.1 | 0.05 | 0.05 | 0.04 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.04 | 0.38 | < 0.1 | < 0.11 | 0.04 | 0.0166 | 0.05 | 0.04 | 0.05 | 0.03 |  |  |  |  | 16
| 1 | 4 | 0.06 | 0.05 | 0.072037 | 0.05 | 0.2 | 0.07 | 0.04 | < 0.1 | 0.04 | 0.05 | 0.04 | 0.03 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.05 | 0.0151 | 0.05 | 0.06 | 0.04 | 0.03 |  |  |  |  | 0.047385
| 1 | 5 | 0.05 | 0.05 | 0.06765 | 0.04 | 0.2 | 0.07 | 0.04 | < 0.1 | 0.05 | 0.05 | 0.05 | 0.04 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.04 | 0.0174 | 0.05 | 0.06 | 0.04 | 0.03 |  |  |  |  | 20
| 1 | 6 | 0.07 | 0.04 | 0.075649 | 0.05 | 0.2 | 0.09 | 0.06 | < 0.1 | 0.04 | 0.05 | 0.05 | 0.05 | < 5 | < 0.05 | < 0.1 | 0.05 | < 0.1 | < 0.1 | < 0.11 | 0.04 | 0.0143 | 0.04 | 0.04 | 0.05 | 0.03
| Mean |  | 0.058333 | 0.045 | 0.071004 | 0.048333 | 0.2 | 0.073333 | 0.048333 |  | 0.045 | 0.048333 | 0.046667 | 0.038333 |  |  |  | 0.048333 | 0.38 |  |  | 0.043333 | 0.0163 | 0.048333 | 0.053333 | 0.046667 | 0.03
| Median |  | 0.06 | 0.045 | 0.070748 | 0.05 | 0.2 | 0.07 | 0.05 |  | 0.045 | 0.05 | 0.05 | 0.04 |  |  |  | 0.05 | 0.38 |  |  | 0.04 | 0.0163 | 0.05 | 0.06 | 0.05 | 0.03
| Std Dev. |  | 0.0075277 | 0.0054772 | 0.0028594 | 0.0040825 | 3.0405e-17 | 0.008165 | 0.0075277 |  | 0.0054772 | 0.0040825 | 0.005164 | 0.0075277 |  |  |  | 0.0040825 |  |  |  | 0.005164 | 0.0014993 | 0.0040825 | 0.010328 | 0.005164 | 0
| Rel.Std.Dev. |  | 12.9% | 12.17% | 4.027% | 8.447% | 1.52e-14% | 11.13% | 15.57% |  | 12.17% | 8.447% | 11.07% | 19.64% |  |  |  | 8.447% |  |  |  | 11.92% | 9.198% | 8.447% | 19.36% | 11.07% | 0%
| PDM3 |  | 23.11% | -5.032% | 49.85% | 2.002% | 322.1% | 54.76% | 2.002% |  | -5.032% | 2.002% | -1.515% | -19.1% |  |  |  | 2.002% | 701.9% |  |  | -8.55% | -65.6% | 2.002% | 12.55% | -1.515% | -36.69%
| Z-Score (Absolute) |  | 1.62 | 0.54 | 3.67 | 0 | NA | 4.05 | 0 | 0.27 | 0.54 | 0 | 0.27 | 1.62 | 396.76 | 3.78 | 0.27 | 0 | 9.17 | 0.27 | 1.08 | 0.81 | 5.18 | 0 | 0.81 | 0.27 | 2.97
Results from laboratory 1.08 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 234b (Certified Value 6.15 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.04 | 6.3817 | 6.2376 | 5.909 | 5.94 | 7.033 | 6.35 | 6.1925 | 5.78 | 6 | 6.46 | 6.55 | 6.5 | 6.08 | 6.12 | 6.337 | 6.02 | 6.373 | 6.14 | 6.0789 | 6.296 | 6.34 | 6.6837 | 6.1 | 5.72 | 5.98 | 6.3 |  |  | 1
| 1 | 2 | 6.02 | 6.4326 | 6.2529 | 6.213 | 6.21 | 7.061 | 6.28 | 6.4207 | 5.72 | 6.1 | 6.11 | 6.36 | 6.12 | 6.12 | 6.06 | 6.414 | 6.01 | 6.138 | 4.84 | 5.8885 | 6.32 | 6.27 | 6.6267 | 5.92 | 5.57 | 5.91 | 6.2 |  |  | #N/A
| 1 | 3 | 5.99 | 6.4564 | 6.2263 | 6.3695 | 6.24 | 6.933 | 6.33 | 6.3203 | 5.73 | 6.21 | 6.28 | 6.21 | 6.27 | 6.16 | 6.09 | 6.406 | 6.02 | 6.104 | 4.8 | 5.8568 | 6.279 | 6.41 | 6.6787 | 5.96 | 5.58 | 5.99 | 6.3 |  |  | 16
| 1 | 4 | 5.98 | 6.5214 | 6.2113 | 6.201 | 5.77 | 6.954 | 6.26 | 6.4267 | 5.74 | 6.14 | 6.14 | 6.14 | 6 | 6.1 | 5.72 | 6.367 | 6 | 6.063 | 4.78 | 5.9361 | 6.254 | 6.35 | 6.4587 | 5.96 | 5.58 | 5.97 | 6.42 |  |  | 6.146
| 1 | 5 | 5.84 | 6.31 | 6.2204 | 5.7774 | 5.68 | 7.004 | 6.39 | 6.3728 | 5.76 | 6.07 | 6.04 | 6.24 | 6.36 | 6.18 | 6.05 | 6.398 | 5.93 | 6.476 | 5.22 | 5.9224 | 6.263 | 6.4 | 6.6727 | 6.05 | 5.68 | 5.91 | 6.65 |  |  | 21
| 1 | 6 | 6.09 | 6.1877 | 6.231 | 6.1877 | 5.78 | 6.99 | 6.22 | 6.2971 | 5.76 | 6.08 | 6.12 | 6.36 | 6.2 | 6.16 | 5.69 | 6.352 | 5.96 | 6.23 | 4.34 | 5.8746 | 6.285 | 6.14 | 6.6797 | 5.94 | 5.56 | 5.95 | 6.4
| Mean |  | 5.9933 | 6.3816 | 6.2299 | 6.1096 | 5.9367 | 6.9958 | 6.305 | 6.3384 | 5.7483 | 6.1 | 6.1917 | 6.31 | 6.2417 | 6.1333 | 5.955 | 6.379 | 5.99 | 6.2307 | 5.02 | 5.9262 | 6.2828 | 6.3183 | 6.6334 | 5.9883 | 5.615 | 5.9517 | 6.3783
| Median |  | 6.005 | 6.4071 | 6.2287 | 6.1944 | 5.86 | 6.997 | 6.305 | 6.3466 | 5.75 | 6.09 | 6.13 | 6.3 | 6.235 | 6.14 | 6.055 | 6.3825 | 6.005 | 6.184 | 4.82 | 5.9055 | 6.282 | 6.345 | 6.6757 | 5.96 | 5.58 | 5.96 | 6.35
| Std Dev. |  | 0.084774 | 0.11871 | 0.014415 | 0.22057 | 0.23872 | 0.047806 | 0.062849 | 0.088401 | 0.022286 | 0.070711 | 0.15316 | 0.14588 | 0.17691 | 0.039328 | 0.19542 | 0.031471 | 0.036878 | 0.16321 | 0.61566 | 0.080411 | 0.023659 | 0.10068 | 0.088127 | 0.070545 | 0.067454 | 0.034881 | 0.15497
| Rel.Std.Dev. |  | 1.414% | 1.86% | 0.2314% | 3.61% | 4.021% | 0.6833% | 0.9968% | 1.395% | 0.3877% | 1.159% | 2.474% | 2.312% | 2.834% | 0.6412% | 3.282% | 0.4933% | 0.6157% | 2.62% | 12.26% | 1.357% | 0.3766% | 1.593% | 1.329% | 1.178% | 1.201% | 0.5861% | 2.43%
| PDM3 |  | -2.484% | 3.834% | 1.366% | -0.5915% | -3.406% | 13.83% | 2.587% | 3.13% | -6.47% | -0.7482% | 0.7433% | 2.669% | 1.557% | -0.2058% | -3.107% | 3.791% | -2.538% | 1.378% | -18.32% | -3.576% | 2.227% | 2.804% | 7.93% | -2.565% | -8.639% | -3.162% | 3.781%
| Z-Score (Absolute) |  | 0.71 | 0.68 | 0.14 | 0.3 | 0.92 | 2.89 | 0.41 | 0.53 | 1.6 | 0.33 | 0 | 0.43 | 0.18 | 0.21 | 0.85 | 0.67 | 0.73 | 0.14 | 4.22 | 0.96 | 0.33 | 0.46 | 1.59 | 0.73 | 2.08 | 0.86 | 0.67
| Analytical results for Cd in OREAS 234b (Certified Value 0.62 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.8 | 0.65 | 0.26 | 0.64372 | 0.4 | 0.7 | 0.68 | 0.58 | 0.9 | 0.7 | 0.62 | 0.66 | 0.57 | 0.65 | < 1 | 0.58 | 0.6 | 0.56 | < 0.1 | 0.5 | 0.57914 | 0.6 | 0.6316 | 0.58 | 0.68 | 0.71 | 0.57 |  |  | 1
| 1 | 2 | 0.71 | 0.7 | 0.26 | 0.6666 | 0.5 | 0.7 | 0.58 | 0.66 | 1 | 0.7 | 0.59 | 0.64 | 0.54 | 0.66 | < 1 | 0.58 | 0.6 | 0.56 | < 0.1 | 0.5 | 0.57956 | 0.58 | 0.6508 | 0.6 | 0.65 | 0.7 | 0.54 |  |  | 24
| 1 | 3 | 0.76 | 0.61 | 0.25 | 0.68266 | 0.5 | 0.8 | 0.6 | 0.62 | 1 | 0.7 | 0.63 | 0.59 | 0.58 | 0.62 | < 1 | 0.57 | 0.6 | 0.61 | < 0.1 | 0.5 | 0.56948 | 0.6 | 0.6418 | 0.57 | 0.64 | 0.68 | 0.57 |  |  | 16
| 1 | 4 | 0.69 | 0.64 | 0.26 | 0.67732 | 0.5 | 0.7 | 0.65 | 0.62 | 1 | 0.7 | 0.59 | 0.61 | 0.6 | 0.59 | < 1 | 0.55 | 0.6 | 0.59 | < 0.1 | 0.5 | 0.56963 | 0.61 | 0.6374 | 0.59 | 0.67 | 0.66 | 0.62 |  |  | 0.61798
| 1 | 5 | 0.71 | 0.61 | 0.27 | 0.64601 | 0.5 | 0.7 | 0.61 | 0.59 | 1 | 0.7 | 0.59 | 0.65 | 0.61 | 0.61 | < 1 | 0.56 | 0.6 | 0.55 | < 0.1 | 0.5 | 0.5813 | 0.62 | 0.6526 | 0.63 | 0.66 | 0.68 | 0.59 |  |  | 22
| 1 | 6 | 0.69 | 0.59 | 0.26 | 0.66802 | 0.5 | 0.8 | 0.66 | 0.65 | 1.3 | 0.7 | 0.61 | 0.63 | 0.59 | 0.65 | < 1 | 0.56 | 0.6 | 0.63 | < 0.1 | 0.5 | 0.57687 | 0.61 | 0.659 | 0.6 | 0.68 | 0.68 | 0.67
| Mean |  | 0.72667 | 0.63333 | 0.26 | 0.66406 | 0.48333 | 0.73333 | 0.63 | 0.62 | 1.0333 | 0.7 | 0.605 | 0.63 | 0.58167 | 0.63 |  | 0.56667 | 0.6 | 0.58333 |  | 0.5 | 0.576 | 0.60333 | 0.64553 | 0.595 | 0.66333 | 0.685 | 0.59333
| Median |  | 0.71 | 0.625 | 0.26 | 0.66731 | 0.5 | 0.7 | 0.63 | 0.62 | 1 | 0.7 | 0.6 | 0.635 | 0.585 | 0.635 |  | 0.565 | 0.6 | 0.575 |  | 0.5 | 0.578 | 0.605 | 0.6463 | 0.595 | 0.665 | 0.68 | 0.58
| Std Dev. |  | 0.044121 | 0.039328 | 0.0063246 | 0.016021 | 0.040825 | 0.05164 | 0.038987 | 0.031623 | 0.13663 | 1.2162e-16 | 0.017607 | 0.026077 | 0.024833 | 0.027568 |  | 0.012111 | 0 | 0.032042 |  | 0 | 0.0051859 | 0.013663 | 0.010327 | 0.020736 | 0.01633 | 0.017607 | 0.045898
| Rel.Std.Dev. |  | 6.072% | 6.21% | 2.433% | 2.413% | 8.447% | 7.042% | 6.188% | 5.1% | 13.22% | 1.737e-14% | 2.91% | 4.139% | 4.269% | 4.376% |  | 2.137% | 0% | 5.493% |  | 0% | 0.9003% | 2.265% | 1.6% | 3.485% | 2.462% | 2.57% | 7.736%
| PDM3 |  | 17.59% | 2.485% | -57.93% | 7.457% | -21.79% | 18.67% | 1.946% | 0.3276% | 67.21% | 13.27% | -2.1% | 1.946% | -5.875% | 1.946% |  | -8.303% | -2.909% | -5.606% |  | -19.09% | -6.793% | -2.369% | 4.459% | -3.718% | 7.34% | 10.85% | -3.988%
| Z-Score (Absolute) |  | 2.9 | 0.67 | 8.21 | 1.41 | NA | NA | 0.59 | 0.36 | NA | NA | 0 | 0.59 | 0.56 | 0.59 | 2.5 | 0.91 | NA | 0.52 | 13.21 | NA | 0.69 | 0.04 | 0.96 | 0.24 | 1.39 | 1.9 | 0.28
Results from laboratories 1.07, 1.08, 1.12, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 234b (Certified Value 13.1 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 12.28 | 13.27 | 11.96 | 12.0206 | 12.7839 | 12.2 | 16 | 12.41 | 13 | 13.4 | 14.25 | 12.3 | 13.15 | < 10 | 12 | 12.2 | 14.5 | 13.2015 | 13.5 | 7.7412 | 15.4 | 13.29 | 14.7 | 12.4 | [5 empty] | 1
| 1 | 2 | 12.52 | 13.68 | 12.34 | 12.0868 | 13.3191 | 13.1 | 16 | 12.85 | 13 | 13.1 | 13.85 | 12.55 | 13.15 | < 10 | 13 | 12.49 | 14.5 | 13.3935 | 13.4 | 7.8253 | 15.4 | 12.85 | 15.09 | 12.2 | [5 empty] | 25
| 1 | 3 | 13.3 | 13.34 | 11.64 | 12.064 | 13.587 | 13.2 | 15 | 12.51 | 13 | 13.35 | 13.7 | 12.55 | 12.5 | < 10 | 12 | 12.14 | 15 | 12.881 | 13.7 | 7.8031 | 15.3 | 12.67 | 14.38 | 12.7 | [5 empty] | 16
| 1 | 4 | 13.52 | 12.76 | 12.12 | 11.9369 | 13.3629 | 12.2 | 16 | 12.88 | 13 | 12.75 | 13.6 | 12.3 | 12.8 | < 10 | 13 | 12.46 | 14 | 12.7905 | 13.3 | 7.4301 | 15.4 | 13.4 | 14.47 | 14.3 | [5 empty] | 13.0911
| 1 | 5 | 12.54 | 13.16 | 12.44 | 11.8939 | 12.7318 | 12.3 | 16 | 12.58 | 13 | 12.9 | 13.9 | 14.3 | 12.85 | < 10 | 12 | 12.23 | 14.5 | 13.2563 | 12.25 | 7.5426 | 15.6 | 13.2 | 14.61 | 13.3 | [5 empty] | 23
| 1 | 6 | 13.71 | 12.94 | 12.21 | 11.977 | 13.4206 | 12.2 | 16 | 12.67 | 13 | 13.25 | 14.15 | 13.45 | 13.15 | < 10 | 12 | 13.04 | 14.5 | 13.3994 | 12.8 | 7.3551 | 15.3 | 13.16 | 14.89 | 15.2
| Mean |  | 12.9783 | 13.1917 | 12.1183 | 11.9965 | 13.2009 | 12.5333 | 15.8333 | 12.65 | 13 | 13.125 | 13.9083 | 12.9083 | 12.9333 |  | 12.3333 | 12.4267 | 14.5 | 13.1537 | 13.1583 | 7.6162 | 15.4 | 13.095 | 14.69 | 13.35
| Median |  | 12.92 | 13.215 | 12.165 | 11.9988 | 13.341 | 12.25 | 16 | 12.625 | 13 | 13.175 | 13.875 | 12.55 | 13 |  | 12 | 12.345 | 14.5 | 13.2289 | 13.35 | 7.6419 | 15.4 | 13.18 | 14.655 | 13
| Std Dev. |  | 0.60367 | 0.32165 | 0.28806 | 0.074513 | 0.35539 | 0.48028 | 0.40825 | 0.1873 | 0 | 0.25836 | 0.25183 | 0.80338 | 0.26583 |  | 0.5164 | 0.33273 | 0.31623 | 0.25961 | 0.53704 | 0.20125 | 0.10954 | 0.27819 | 0.26495 | 1.1811
| Rel.Std.Dev. |  | 4.651% | 2.438% | 2.377% | 0.6211% | 2.692% | 3.832% | 2.578% | 1.481% | 0% | 1.968% | 1.811% | 6.224% | 2.055% |  | 4.187% | 2.678% | 2.181% | 1.974% | 4.081% | 2.642% | 0.7113% | 2.124% | 1.804% | 8.847%
| PDM3 |  | -0.8612% | 0.7684% | -7.431% | -8.361% | 0.8387% | -4.261% | 20.95% | -3.369% | -0.6957% | 0.2591% | 6.243% | -1.396% | -1.205% |  | -5.788% | -5.075% | 10.76% | 0.4784% | 0.5137% | -41.82% | 17.64% | 0.02994% | 12.21% | 1.978%
| Z-Score (Absolute) |  | 0.18 | 0.15 | 1.48 | 1.66 | 0.16 | 0.85 | NA | 0.67 | NA | 0.05 | 1.23 | 0.28 | 0.24 | 12.27 | NA | 1.01 | 2.13 | 0.09 | 0.1 | 8.3 | 3.49 | 0 | 2.42 | 0.39
Results from laboratories 1.08, 1.13 and 1.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 234b (Certified Value 43.2 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 47.6 | 44.4 | 42.22 | 35.6762 | 40.6092 | 42 | 52 | 44 | 42.2 | 33 | 42.6 | 43.6 | 46.7 | 42.9 | 44 | 43 | 41.3 | 41.7 | 43.7 | 44.78 | 40 | 43.6888 | 43.1 | 42.9677 | 50.3 | 46.4 | 49.2 | 40.3 |  | 1
| 1 | 2 | 47.9 | 44.4 | 41.78 | 36.4708 | 43.4305 | 42 | 48.9 | 44 | 44.9 | 32 | 41.8 | 43.6 | 46.7 | 42.9 | 44 | 43 | 41.8 | 42.4 | 43.5 | 37.39 | 40 | 43.8321 | 42.8 | 42.4238 | 48.8 | 44.3 | 50.3 | 39.7 |  | 26
| 1 | 3 | 46.2 | 43.7 | 41.17 | 36.3404 | 44.2526 | 42 | 48.8 | 44 | 43.7 | 33 | 43.2 | 45.4 | 46.3 | 43.1 | 42 | 43 | 40.8 | 42.2 | 44.2 | 37.84 | 40 | 43.9733 | 44.1 | 42.1745 | 49.7 | 45.4 | 47.8 | 41.2 |  | 16
| 1 | 4 | 47.9 | 44 | 41.2 | 35.5263 | 42.717 | 43 | 50.2 | 43 | 44.4 | 33 | 43.3 | 43.7 | 47.3 | 41.7 | 42.9 | 41 | 40.6 | 41 | 44.7 | 36.94 | 40 | 45.9531 | 43.9 | 40.6235 | 50.7 | 45.6 | 48.1 | 43.5 |  | 43.195
| 1 | 5 | 46.6 | 43 | 41.15 | 36.0651 | 39.8061 | 42 | 50.5 | 44 | 43.8 | 33 | 43 | 44.4 | 45 | 47.2 | 43.8 | 43 | 41.5 | 41.3 | 43.3 | 37.51 | 40 | 44.5611 | 43.8 | 40.8939 | 50.1 | 44 | 48.2 | 40.8 |  | 24
| 1 | 6 | 47.6 | 42.5 | 41.89 | 36.149 | 42.7394 | 42 | 50.3 | 44 | 43.2 | 32 | 42 | 44.1 | 46.3 | 47.7 | 44.9 | 40 | 42 | 40.9 | 46 | 35.77 | 40 | 44.6784 | 41.8 | 41.0084 | 50.5 | 45.1 | 48.7 | 46.3
| Mean |  | 47.3 | 43.6667 | 41.5683 | 36.038 | 42.2591 | 42.1667 | 50.1167 | 43.8333 | 43.7 | 32.6667 | 42.65 | 44.1333 | 46.3833 | 44.25 | 43.6 | 42.1667 | 41.3333 | 41.5833 | 44.2333 | 38.3717 | 40 | 44.4478 | 43.25 | 41.682 | 50.0167 | 45.1333 | 48.7167 | 41.9667
| Median |  | 47.6 | 43.85 | 41.49 | 36.107 | 42.7282 | 42 | 50.25 | 44 | 43.75 | 33 | 42.8 | 43.9 | 46.5 | 43 | 43.9 | 43 | 41.4 | 41.5 | 43.95 | 37.45 | 40 | 44.2672 | 43.45 | 41.5915 | 50.2 | 45.25 | 48.45 | 41
| Std Dev. |  | 0.72111 | 0.77374 | 0.45657 | 0.37002 | 1.7045 | 0.40825 | 1.1788 | 0.40825 | 0.94234 | 0.5164 | 0.63166 | 0.69761 | 0.7705 | 2.5328 | 1.01 | 1.3292 | 0.55015 | 0.62423 | 1.0033 | 3.221 | 0 | 0.83813 | 0.86891 | 0.96344 | 0.68823 | 0.88015 | 0.9196 | 2.4897
| Rel.Std.Dev. |  | 1.525% | 1.772% | 1.098% | 1.027% | 4.033% | 0.9682% | 2.352% | 0.9314% | 2.156% | 1.581% | 1.481% | 1.581% | 1.661% | 5.724% | 2.316% | 3.152% | 1.331% | 1.501% | 2.268% | 8.394% | 0% | 1.886% | 2.009% | 2.311% | 1.376% | 1.95% | 1.888% | 5.933%
| PDM3 |  | 9.503% | 1.092% | -3.766% | -16.57% | -2.167% | -2.381% | 16.02% | 1.478% | 1.169% | -24.37% | -1.262% | 2.172% | 7.381% | 2.442% | 0.9376% | -2.381% | -4.31% | -3.731% | 2.404% | -11.17% | -7.397% | 2.9% | 0.1273% | -3.503% | 15.79% | 4.487% | 12.78% | -2.844%
| Z-Score (Absolute) |  | 1.65 | 0.1 | 0.79 | 3.15 | 0.5 | 0.54 | 2.85 | 0.17 | 0.12 | 4.58 | 0.33 | 0.3 | 1.26 | 0.35 | 0.07 | 0.54 | 0.89 | 0.78 | 0.34 | 2.15 | 1.46 | 0.44 | 0.07 | 0.74 | 2.81 | 0.73 | 2.25 | 0.62
| Analytical results for Cr in OREAS 234b (Certified Value 144 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 110 | 174 | 127.7632 | 104 | 207 | 148 | 123 | 150 | 164 | 142 | 151 | 147 | 130 | 151 | 126 | 145 | 153 | 181.15 | 163.467 | 165 | 171.0065 | 140 | 154.182 | 140 | 158 | 157 | 122 |  |  | 1
| 1 | 2 | 109 | 165 | 136.461 | 113 | 214 | 137 | 135 | 151 | 161 | 134 | 156 | 140 | 132 | 148 | 123 | 144 | 144 | 148.73 | 163.242 | 160 | 173.8225 | 139 | 152.6849 | 135 | 153 | 155 | 122 |  |  | 27
| 1 | 3 | 119 | 179 | 136.9305 | 111 | 207 | 144 | 124 | 151 | 163 | 136 | 149 | 138 | 131 | 149 | 128 | 144 | 148 | 147.78 | 139.536 | 165 | 172.3369 | 143 | 156.2897 | 142 | 154 | 159 | 122 |  |  | 16
| 1 | 4 | 119 | 177 | 133.9266 | 102 | 206 | 139 | 154 | 153 | 157 | 133 | 150 | 134 | 129 | 144 | 124 | 147 | 148 | 145.21 | 144.27 | 155 | 177.4065 | 142 | 149.1891 | 142 | 154 | 161 | 147 |  |  | 144.1065
| 1 | 5 | 115 | 167 | 125.8034 | 103 | 212 | 139 | 110 | 154 | 167 | 131 | 150 | 141 | 128 | 148 | 126 | 142 | 154 | 149.98 | 155.178 | 160 | 173.0549 | 140 | 154.5805 | 149 | 155 | 161 | 152 |  |  | 25
| 1 | 6 | 112 | 162 | 134.3043 | 115 | 209 | 140 | 120 | 153 | 163 | 131 | 145 | 138 | 131 | 138 | 121 | 142 | 149 | 144.32 | 153.648 | 165 | 176.6202 | 141 | 156.28 | 147 | 155 | 164 | 147
| Mean |  | 114 | 170.6667 | 132.5315 | 108 | 209.1667 | 141.1667 | 127.6667 | 152 | 162.5 | 134.5 | 150.1667 | 139.6667 | 130.1667 | 146.3333 | 124.6667 | 144 | 149.3333 | 152.8617 | 153.2235 | 161.6667 | 174.0412 | 140.8333 | 153.8677 | 142.5 | 154.8333 | 159.5 | 135.3333
| Median |  | 113.5 | 170.5 | 134.1155 | 107.5 | 208 | 139.5 | 123.5 | 152 | 163 | 133.5 | 150 | 139 | 130.5 | 148 | 125 | 144 | 148.5 | 148.255 | 154.413 | 162.5 | 173.4387 | 140.5 | 154.3812 | 142 | 154.5 | 160 | 134.5
| Std Dev. |  | 4.3818 | 6.9474 | 4.6451 | 5.6569 | 3.1885 | 4.0702 | 15.1877 | 1.5492 | 3.3317 | 4.1352 | 3.5449 | 4.3205 | 1.472 | 4.6762 | 2.5033 | 1.8974 | 3.6697 | 14.021 | 9.7654 | 4.0825 | 2.4944 | 1.472 | 2.6676 | 5.01 | 1.7224 | 3.2094 | 14.7196
| Rel.Std.Dev. |  | 3.844% | 4.071% | 3.505% | 5.238% | 1.524% | 2.883% | 11.9% | 1.019% | 2.05% | 3.075% | 2.361% | 3.093% | 1.131% | 3.196% | 2.008% | 1.318% | 2.457% | 9.172% | 6.373% | 2.525% | 1.433% | 1.045% | 1.734% | 3.516% | 1.112% | 2.012% | 10.88%
| PDM3 |  | -20.89% | 18.43% | -8.032% | -25.06% | 45.15% | -2.04% | -11.41% | 5.478% | 12.76% | -6.666% | 4.205% | -3.081% | -9.673% | 1.545% | -13.49% | -0.07388% | 3.627% | 6.076% | 6.327% | 12.19% | 20.77% | -2.271% | 6.774% | -1.115% | 7.444% | 10.68% | -6.088%
| Z-Score (Absolute) |  | 1.98 | 1.49 | 0.85 | 2.35 | 3.85 | 0.32 | 1.14 | 0.35 | 0.99 | 0.73 | 0.23 | 0.41 | 0.99 | 0 | 1.33 | 0.14 | 0.18 | 0.4 | 0.42 | 0.94 | 1.7 | 0.34 | 0.46 | 0.23 | 0.52 | 0.81 | 0.67
| Analytical results for Cs in OREAS 234b (Certified Value 0.95 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | < 5 | 0.96 | 0.91194 | 0.95 | 1 | 0.9 | 0.93 | 1 | 0.96 | 1 | 0.95 | 0.97 | 0.9 | 0.92 | 1 | 0.96297 | 0.98 | 1.0303 | 0.92 | 1.13 | 1.05 | 0.86 | [7 empty] | 1
| 1 | 2 | < 5 | 0.99 | 0.89307 | 0.92 | 1 | 0.9 | 1.01 | 1.1 | 0.95 | 0.98 | 0.94 | 0.97 | 1 | 0.9 | 1 | 0.954 | 0.99 | 1.0142 | 0.89 | 1.01 | 1.05 | 0.85 | [7 empty] | 28
| 1 | 3 | < 5 | 0.95 | 0.85213 | 0.92 | 1 | 1 | 0.99 | 1.1 | 0.96 | 0.92 | 0.96 | 0.92 | 0.9 | 0.93 | 1 | 0.94522 | 0.99 | 1.0324 | 0.86 | 1.09 | 1 | 0.89 | [7 empty] | 16
| 1 | 4 | < 5 | 0.97 | 0.91641 | 0.89 | 1 | 0.9 | 1 | 1.1 | 0.94 | 0.98 | 0.94 | 0.94 | 0.9 | 0.92 | 1 | 0.94153 | 0.98 | 0.9743 | 0.88 | 1.16 | 1.01 | 0.98 | [7 empty] | 0.95225
| 1 | 5 | < 5 | 0.96 | 0.87868 | 0.89 | 1 | 0.9 | 0.98 | 1 | 0.92 | 0.98 | 1.02 | 0.94 | 0.9 | 0.87 | 1 | 0.95747 | 0.9 | 0.9657 | 0.89 | 1.13 | 1.03 | 0.89 | [7 empty] | 26
| 1 | 6 | < 5 | 0.93 | 0.91353 | 0.86 | 1 | 0.9 | 0.93 | 1.1 | 0.93 | 1 | 1.03 | 0.97 | 1 | 0.96 | 1 | 0.98504 | 0.96 | 0.9752 | 0.9 | 1.15 | 1.08 | 1.02
| Mean |  |  | 0.96 | 0.89429 | 0.905 | 1 | 0.91667 | 0.97333 | 1.0667 | 0.94333 | 0.97667 | 0.97333 | 0.95167 | 0.93333 | 0.91667 | 1 | 0.9577 | 0.96667 | 0.99868 | 0.89 | 1.1117 | 1.0367 | 0.915
| Median |  |  | 0.96 | 0.90251 | 0.905 | 1 | 0.9 | 0.985 | 1.1 | 0.945 | 0.98 | 0.955 | 0.955 | 0.9 | 0.92 | 1 | 0.95573 | 0.98 | 0.9947 | 0.89 | 1.13 | 1.04 | 0.89
| Std Dev. |  |  | 0.02 | 0.025273 | 0.031464 | 0 | 0.040825 | 0.035024 | 0.05164 | 0.01633 | 0.029439 | 0.040825 | 0.02137 | 0.05164 | 0.030111 | 0 | 0.015532 | 0.034448 | 0.030368 | 0.02 | 0.055287 | 0.029439 | 0.06892
| Rel.Std.Dev. |  |  | 2.083% | 2.826% | 3.477% | 0% | 4.454% | 3.598% | 4.841% | 1.731% | 3.014% | 4.194% | 2.246% | 5.533% | 3.285% | 0% | 1.622% | 3.564% | 3.041% | 2.247% | 4.973% | 2.84% | 7.532%
| PDM3 |  |  | 0.8143% | -6.086% | -4.961% | 5.015% | -3.736% | 2.215% | 12.02% | -0.9359% | 2.565% | 2.215% | -0.06078% | -1.986% | -3.736% | 5.015% | 0.5733% | 1.514% | 4.877% | -6.537% | 16.74% | 8.865% | -3.911%
| Z-Score (Absolute) |  | 26.84 | 0 | 1.15 | 0.96 | NA | NA | 0.23 | NA | 0.29 | 0.29 | 0.23 | 0.15 | NA | 0.76 | NA | 0.04 | 0.12 | 0.67 | 1.22 | 2.64 | 1.34 | 0.78
| Analytical results for Cu in OREAS 234b (Certified Value 162 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 168.3 | 160.4 | 151.92 | 174.14 | 152.232 | 164 | 179 | 159 | 152.5 | 161 | 160.7 | 167 | 169 | 167 | 168.5 | 174 | 162 | 148.5 | 170 | 27.51 | 160.018 | 155 | 150.3958 | 162 | 133.1008 | 178 | 164.7 | 159.5 | 158 | 1
| 1 | 2 | 171 | 164.6 | 152.09 | 175.21 | 160.6462 | 166 | 169 | 156 | 162.7 | 159 | 151.9 | 163.5 | 164.5 | 158.5 | 166.5 | 172 | 163 | 146.7 | 164 | 23.59 | 157.239 | 155 | 152.983 | 160 | 130.4418 | 172 | 161.5 | 160.2 | 151 | 29
| 1 | 3 | 167.2 | 161.3 | 150.04 | 174.01 | 165.2953 | 165 | 173 | 158 | 156.5 | 158 | 162.7 | 167.5 | 165 | 161.5 | 166.5 | 173 | 163 | 149.2 | 162 | 27.86 | 152.955 | 155 | 154.3855 | 160.5 | 131.049 | 175 | 165.1 | 163.4 | 155 | 16
| 1 | 4 | 163 | 160.5 | 149.4 | 174.2 | 161.173 | 166 | 177 | 157 | 157.6 | 161 | 162.2 | 162 | 163.5 | 158 | 162.5 | 163 | 162 | 145.4 | 167 | 23.52 | 156.789 | 155 | 162.665 | 161.5 | 124.5073 | 173 | 158.2 | 162.7 | 151 | 161.8509
| 1 | 5 | 166.1 | 157.9 | 152.36 | 175.06 | 149.6235 | 166 | 178 | 160 | 159.2 | 162 | 158.7 | 160.5 | 163 | 168.5 | 167.5 | 171 | 163 | 147.3 | 171 | 25.41 | 158.058 | 155 | 156.3927 | 162.5 | 123.3251 | 179 | 160.4 | 161.5 | 155 | 27
| 1 | 6 | 169.9 | 155 | 151.12 | 176.51 | 160.6543 | 165 | 174 | 153 | 156.9 | 158 | 156.5 | 163 | 166.5 | 167.5 | 167 | 159 | 162 | 145.8 | 169 | 25.06 | 154.56 | 155 | 155.2151 | 158.5 | 122.0673 | 174 | 160.8 | 162.8 | 151
| Mean |  | 167.5833 | 159.95 | 151.155 | 174.855 | 158.2707 | 165.3333 | 175 | 157.1667 | 157.5667 | 159.8333 | 158.7833 | 163.9167 | 165.25 | 163.5 | 166.4167 | 168.6667 | 162.5 | 147.15 | 167.1667 | 25.4917 | 156.6032 | 155 | 155.3395 | 160.8333 | 127.4152 | 175.1667 | 161.7833 | 161.6833 | 153.5
| Median |  | 167.75 | 160.45 | 151.52 | 174.63 | 160.6503 | 165.5 | 175.5 | 157.5 | 157.25 | 160 | 159.7 | 163.25 | 164.75 | 164.25 | 166.75 | 171.5 | 162.5 | 147 | 168 | 25.235 | 157.014 | 155 | 154.8003 | 161 | 127.4746 | 174.5 | 161.15 | 162.1 | 153
| Std Dev. |  | 2.8604 | 3.2452 | 1.203 | 0.95513 | 6.0055 | 0.8165 | 3.7417 | 2.4833 | 3.3548 | 1.7224 | 4.0814 | 2.7824 | 2.2079 | 4.7434 | 2.0595 | 6.1536 | 0.54772 | 1.4923 | 3.5449 | 1.8644 | 2.5179 | 0 | 4.1398 | 1.472 | 4.6578 | 2.7869 | 2.6574 | 1.5639 | 2.9496
| Rel.Std.Dev. |  | 1.707% | 2.029% | 0.7959% | 0.5462% | 3.794% | 0.4938% | 2.138% | 1.58% | 2.129% | 1.078% | 2.57% | 1.697% | 1.336% | 2.901% | 1.238% | 3.648% | 0.3371% | 1.014% | 2.121% | 7.314% | 1.608% | 0% | 2.665% | 0.9152% | 3.656% | 1.591% | 1.643% | 0.9672% | 1.922%
| PDM3 |  | 3.542% | -1.174% | -6.608% | 8.035% | -2.212% | 2.152% | 8.124% | -2.894% | -2.647% | -1.247% | -1.895% | 1.276% | 2.1% | 1.019% | 2.821% | 4.211% | 0.4011% | -9.083% | 3.284% | -84.25% | -3.242% | -4.233% | -4.023% | -0.6287% | -21.28% | 8.227% | -0.04173% | -0.1035% | -5.16%
| Z-Score (Absolute) |  | 1.01 | 0.13 | 1.45 | 2.1 | 0.38 | 0.67 | 2.12 | 0.55 | 0.49 | 0.15 | 0.31 | 0.46 | 0.66 | 0.4 | 0.84 | 1.17 | 0.25 | 2.05 | 0.95 | 20.28 | 0.63 | 0.87 | 0.82 | 0 | 5.01 | 2.15 | 0.14 | 0.13 | 1.1
| Analytical results for Dy in OREAS 234b (Certified Value 3.59 ppm) | [30 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.71 | 3.22 | 3.283 | 3.4 | 3.74 | 3.51 | 3.8 | 3.639 | [21 empty] | 1
| 1 | 2 | 3.68 | 3.37 | 3.2623 | 3.7 | 3.56 | 3.62 | 3.8 | 3.6241 | [21 empty] | 30
| 1 | 3 | 3.65 | 3.2 | 3.2807 | 3.6 | 3.56 | 3.51 | 3.6 | 3.5877 | [21 empty] | 16
| 1 | 4 | 3.68 | 3.33 | 3.2789 | 3.3 | 3.77 | 3.58 | 3.8 | 3.544 | [21 empty] | 3.5898
| 1 | 5 | 3.65 | 3.48 | 3.2548 | 3.5 | 3.67 | 3.42 | 3.9 | 3.7322 | [21 empty] | 28
| 1 | 6 | 3.57 | 3.34 | 3.2904 | 3.6 | 3.71 | 3.64 | 3.7 | 3.774
| Mean |  | 3.6567 | 3.3233 | 3.275 | 3.5167 | 3.6683 | 3.5467 | 3.7667 | 3.6502
| Median |  | 3.665 | 3.335 | 3.2798 | 3.55 | 3.69 | 3.545 | 3.8 | 3.6316
| Std Dev. |  | 0.048028 | 0.10289 | 0.013547 | 0.1472 | 0.090203 | 0.082381 | 0.10328 | 0.087219
| Rel.Std.Dev. |  | 1.313% | 3.096% | 0.4136% | 4.186% | 2.459% | 2.323% | 2.742% | 2.389%
| PDM3 |  | 1.863% | -7.422% | -8.769% | -2.037% | 2.188% | -1.201% | 4.927% | 1.682%
| Z-Score (Absolute) |  | 0.52 | 2.45 | 2.88 | 0.73 | 0.62 | 0.46 | 1.5 | 0.46
| Analytical results for Er in OREAS 234b (Certified Value 2.17 ppm) | [30 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.25 | 1.9 | 2.0359 | 2.2 | 2.29 | 2.1 | 2.4 | 2.3367 | [21 empty] | 1
| 1 | 2 | 2.27 | 2.02 | 2.0419 | 2.3 | 2.18 | 2.1 | 2.2 | 2.3547 | [21 empty] | 31
| 1 | 3 | 2.27 | 1.9 | 2.0974 | 2.1 | 2.2 | 2.11 | 2.1 | 2.2821 | [21 empty] | 16
| 1 | 4 | 2.2 | 1.96 | 2.0428 | 2.1 | 2.3 | 2.16 | 2.2 | 2.2097 | [21 empty] | 2.1733
| 1 | 5 | 2.27 | 2.06 | 2.0802 | 2.2 | 2.26 | 2.04 | 2.3 | 2.3588 | [21 empty] | 29
| 1 | 6 | 2.23 | 1.98 | 2.066 | 2.1 | 2.23 | 2.13 | 2.4 | 2.4033
| Mean |  | 2.2483 | 1.97 | 2.0607 | 2.1667 | 2.2433 | 2.1067 | 2.2667 | 2.3242
| Median |  | 2.26 | 1.97 | 2.0544 | 2.15 | 2.245 | 2.105 | 2.25 | 2.3457
| Std Dev. |  | 0.028577 | 0.064187 | 0.024684 | 0.08165 | 0.048442 | 0.039833 | 0.12111 | 0.068448
| Rel.Std.Dev. |  | 1.271% | 3.258% | 1.198% | 3.768% | 2.159% | 1.891% | 5.343% | 2.945%
| PDM3 |  | 3.451% | -9.356% | -5.182% | -0.3064% | 3.221% | -3.067% | 4.295% | 6.944%
| Z-Score (Absolute) |  | 0.37 | 1.98 | 1.22 | 0.32 | 0.32 | 0.83 | 0.52 | 1
| Analytical results for Eu in OREAS 234b (Certified Value 0.91 ppm) | [30 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.92 | 0.84 | 1.3035 | 0.84 | 0.96 | 0.92 | 1 | 0.88486 | [21 empty] | 1
| 1 | 2 | 0.93 | 0.87 | 1.2709 | 0.89 | 0.92 | 0.94 | 0.9 | 0.91216 | [21 empty] | 32
| 1 | 3 | 0.94 | 0.83 | 1.2745 | 0.88 | 0.91 | 0.94 | 1 | 0.91527 | [21 empty] | 16
| 1 | 4 | 0.88 | 0.86 | 1.2799 | 0.81 | 0.96 | 0.97 | 1 | 0.88402 | [21 empty] | 0.91333
| 1 | 5 | 0.91 | 0.9 | 1.2819 | 0.85 | 0.94 | 0.93 | 1 | 0.8904 | [21 empty] | 30
| 1 | 6 | 0.87 | 0.85 | 1.253 | 0.86 | 0.94 | 0.97 | 1 | 0.94298
| Mean |  | 0.90833 | 0.85833 | 1.2773 | 0.855 | 0.93833 | 0.945 | 0.98333 | 0.90495
| Median |  | 0.915 | 0.855 | 1.2772 | 0.855 | 0.94 | 0.94 | 1 | 0.90128
| Std Dev. |  | 0.027869 | 0.024833 | 0.016449 | 0.02881 | 0.020412 | 0.020736 | 0.040825 | 0.023057
| Rel.Std.Dev. |  | 3.068% | 2.893% | 1.288% | 3.37% | 2.175% | 2.194% | 4.152% | 2.548%
| PDM3 |  | -0.5466% | -6.021% | 39.85% | -6.386% | 2.738% | 3.468% | 7.665% | -0.9173%
| Z-Score (Absolute) |  | 0.25 | 1.07 | 5.85 | 1.13 | 0.25 | 0.36 | 0.99 | 0.3
| Analytical results for Fe in OREAS 234b (Certified Value 7.83 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 8.06 | 8.09 | 8.0091 | 7.9571 | 7.38 | 8.883 | 8.19 | 7.55 | 7.38 | 7.4 | 8.21 | 8.08 | 8.32 | 7.95 | 7.66 | 8.2 | 7.39 | 7.98 | 10.17 | 7.6309 | 7.945 | 7.92 | 8.0533 | 7.84 | 7.38 | 8.33 | 8.4 |  |  | 1
| 1 | 2 | 8.14 | 8.08 | 7.9838 | 7.8862 | 7.75 | 8.883 | 8.04 | 7.88 | 7.35 | 7.46 | 8.02 | 7.88 | 7.73 | 7.91 | 7.57 | 8.4 | 7.39 | 7.65 | 7.91 | 7.3822 | 7.95 | 7.86 | 7.9423 | 7.6 | 7.34 | 8.28 | 8.3 |  |  | #N/A
| 1 | 3 | 8.06 | 8.15 | 8.0309 | 8.0916 | 7.77 | 8.603 | 8.17 | 7.8 | 7.39 | 7.71 | 8.23 | 7.68 | 7.98 | 7.98 | 7.49 | 8.31 | 7.36 | 7.63 | 7.91 | 7.4075 | 8.005 | 7.92 | 8.0443 | 7.62 | 7.36 | 8.35 | 8.5 |  |  | 16
| 1 | 4 | 8.03 | 8.09 | 7.9716 | 7.8082 | 7.27 | 8.673 | 8.04 | 8 | 7.42 | 7.61 | 8.02 | 7.65 | 7.63 | 7.91 | 7.14 | 8.25 | 7.34 | 7.64 | 7.51 | 7.4594 | 7.91 | 7.9 | 7.7503 | 7.64 | 7.25 | 8.32 | 8.2 |  |  | 7.8271
| 1 | 5 | 7.89 | 7.93 | 8.0045 | 7.8064 | 7.14 | 8.673 | 8.24 | 7.85 | 7.4 | 7.59 | 7.96 | 7.68 | 8.24 | 8.11 | 7.5 | 8.42 | 7.31 | 8.12 | 7.94 | 7.5015 | 7.945 | 7.98 | 8.0383 | 7.79 | 7.31 | 8.26 | 7.3 |  |  | 31
| 1 | 6 | 8.18 | 8.09 | 8.0767 | 7.7725 | 7.09 | 8.603 | 8.04 | 7.69 | 7.36 | 7.51 | 8.09 | 7.86 | 8.01 | 8.06 | 7 | 8.15 | 7.3 | 7.83 | 7.56 | 7.3059 | 7.98 | 7.55 | 7.9373 | 7.61 | 7.44 | 8.31 | 9
| Mean |  | 8.06 | 8.0717 | 8.0128 | 7.887 | 7.4 | 8.7197 | 8.12 | 7.795 | 7.3833 | 7.5467 | 8.0883 | 7.805 | 7.985 | 7.9867 | 7.3933 | 8.2883 | 7.3483 | 7.8083 | 8.1667 | 7.4479 | 7.9558 | 7.855 | 7.961 | 7.6833 | 7.3467 | 8.3083 | 8.2833
| Median |  | 8.06 | 8.09 | 8.0068 | 7.8472 | 7.325 | 8.673 | 8.105 | 7.825 | 7.385 | 7.55 | 8.055 | 7.77 | 7.995 | 7.965 | 7.495 | 8.28 | 7.35 | 7.74 | 7.91 | 7.4335 | 7.9475 | 7.91 | 7.9903 | 7.63 | 7.35 | 8.315 | 8.35
| Std Dev. |  | 0.1006 | 0.073869 | 0.03751 | 0.12048 | 0.29678 | 0.13033 | 0.090554 | 0.15707 | 0.02582 | 0.11219 | 0.11017 | 0.16706 | 0.27164 | 0.082138 | 0.26151 | 0.10834 | 0.038687 | 0.20605 | 0.99957 | 0.1119 | 0.032774 | 0.15437 | 0.11552 | 0.10405 | 0.064395 | 0.033116 | 0.55648
| Rel.Std.Dev. |  | 1.248% | 0.9152% | 0.4681% | 1.528% | 4.011% | 1.495% | 1.115% | 2.015% | 0.3497% | 1.487% | 1.362% | 2.14% | 3.402% | 1.028% | 3.537% | 1.307% | 0.5265% | 2.639% | 12.24% | 1.502% | 0.412% | 1.965% | 1.451% | 1.354% | 0.8765% | 0.3986% | 6.718%
| PDM3 |  | 2.975% | 3.125% | 2.372% | 0.7652% | -5.457% | 11.4% | 3.742% | -0.4102% | -5.67% | -3.583% | 3.337% | -0.2824% | 2.017% | 2.039% | -5.542% | 5.893% | -6.117% | -0.2399% | 4.338% | -4.845% | 1.645% | 0.3564% | 1.71% | -1.837% | -6.138% | 6.148% | 5.829%
| Z-Score (Absolute) |  | 0.43 | 0.48 | 0.23 | 0.28 | 2.28 | 3.14 | 0.67 | 0.66 | 2.35 | 1.68 | 0.54 | 0.62 | 0.12 | 0.13 | 2.31 | 1.37 | 2.5 | 0.61 | 0.87 | 2.09 | 0 | 0.41 | 0.02 | 1.12 | 2.5 | 1.45 | 1.35
| Analytical results for Ga in OREAS 234b (Certified Value 15.2 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 16.6 | 15.04 | 14.2836 | 13.0829 | 13.3 | 15.1 | 15.7 | 15.24 | 18 | 15 | 15.5 | 16.9 | 14.9 | 15.6 | 11 | 13.7 | 14.84 | 14.41 | 12 | 15.1103 | 15.85 | 16.6585 | 18.7 | 16 | 16.11 | 13.5 |  |  |  | 1
| 1 | 2 | 16.9 | 15.43 | 14.2655 | 13.6442 | 13.7 | 14.4 | 15.1 | 15.89 | 18 | 15 | 14.65 | 15.85 | 14.9 | 15.45 | 11 | 15.1 | 15.41 | 10.3 | 11 | 15.2123 | 15.75 | 16.5598 | 18.1 | 15.2 | 15.8 | 13.5 |  |  |  | 33
| 1 | 3 | 16.8 | 15.73 | 14.0931 | 13.8937 | 14 | 14.6 | 15.1 | 15.48 | 18 | 15 | 15 | 16.1 | 14.9 | 15.2 | 11 | 14.2 | 15.76 | 13.14 | 11 | 15.0936 | 15.55 | 16.7558 | 18.3 | 15.5 | 15.48 | 14 |  |  |  | 16
| 1 | 4 | 17 | 15.41 | 14.2956 | 13.7394 | 12.8 | 14.7 | 15.3 | 15.91 | 18 | 15 | 14.55 | 16.3 | 14.75 | 15.15 | 10 | 13.9 | 15.77 | 12.06 | 12 | 15.2685 | 15.75 | 16.3088 | 18.5 | 16 | 15.21 | 15.5 |  |  |  | 15.239
| 1 | 5 | 16.8 | 15.36 | 14.221 | 12.6793 | 12.8 | 14.9 | 15.3 | 15.98 | 19 | 15 | 14.6 | 16.5 | 15.6 | 15.05 | 11 | 14.8 | 14.64 | 12.8 | 12 | 15.3695 | 14.35 | 16.6005 | 18.4 | 15.3 | 15.49 | 14.4 |  |  |  | 32
| 1 | 6 | 17.3 | 14.68 | 14.452 | 13.5977 | 12.4 | 14.7 | 15.1 | 15.4 | 18 | 15 | 14.9 | 16.35 | 15.95 | 15.4 | 10 | 14.9 | 15.48 | 12.91 | 12 | 15.1402 | 14.8 | 16.6372 | 18.1 | 15.6 | 15.29 | 16.4
| Mean |  | 16.9 | 15.275 | 14.2685 | 13.4395 | 13.1667 | 14.7333 | 15.2667 | 15.65 | 18.1667 | 15 | 14.8667 | 16.3333 | 15.1667 | 15.3083 | 10.6667 | 14.4333 | 15.3167 | 12.6033 | 11.6667 | 15.199 | 15.3417 | 16.5868 | 18.35 | 15.6 | 15.5633 | 14.55
| Median |  | 16.85 | 15.385 | 14.2745 | 13.621 | 13.05 | 14.7 | 15.2 | 15.685 | 18 | 15 | 14.775 | 16.325 | 14.9 | 15.3 | 11 | 14.5 | 15.445 | 12.855 | 12 | 15.1762 | 15.65 | 16.6189 | 18.35 | 15.55 | 15.485 | 14.2
| Std Dev. |  | 0.23664 | 0.3649 | 0.11632 | 0.46223 | 0.60882 | 0.24221 | 0.23381 | 0.3142 | 0.40825 | 0 | 0.3573 | 0.3573 | 0.48751 | 0.20837 | 0.5164 | 0.5785 | 0.47382 | 1.3627 | 0.5164 | 0.10631 | 0.6184 | 0.15129 | 0.23452 | 0.34059 | 0.33667 | 1.1709
| Rel.Std.Dev. |  | 1.4% | 2.389% | 0.8152% | 3.439% | 4.624% | 1.644% | 1.532% | 2.008% | 2.247% | 0% | 2.403% | 2.188% | 3.214% | 1.361% | 4.841% | 4.008% | 3.094% | 10.81% | 4.426% | 0.6994% | 4.031% | 0.9121% | 1.278% | 2.183% | 2.163% | 8.047%
| PDM3 |  | 10.9% | 0.236% | -6.369% | -11.81% | -13.6% | -3.318% | 0.1813% | 2.697% | 19.21% | -1.569% | -2.444% | 7.181% | -0.4749% | 0.4548% | -30% | -5.287% | 0.5094% | -17.3% | -23.44% | -0.2624% | 0.6735% | 8.844% | 20.41% | 2.369% | 2.128% | -4.522%
| Z-Score (Absolute) |  | 2.75 | 0.01 | 1.68 | 3.08 | 3.54 | 0.9 | 0 | 0.65 | NA | 0.45 | 0.67 | 1.8 | 0.17 | 0.07 | NA | 1.4 | 0.08 | 4.49 | NA | 0.11 | 0.13 | 2.23 | 5.2 | 0.56 | 0.5 | 1.21
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 234b (Certified Value 2.96 ppm) | [30 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.19 | 2.64 | 3.2427 | 2.8 | 2.8 | 3.1 | 2.9 | 0.54675 | [21 empty] | 1
| 1 | 2 | 3.35 | 2.77 | 3.2681 | 3 | 2.7 | 3.1 | 2.8 | 0.56684 | [21 empty] | 34
| 1 | 3 | 3.2 | 2.62 | 3.2379 | 2.9 | 2.7 | 3.07 | 2.9 | 0.54357 | [21 empty] | 16
| 1 | 4 | 3.1 | 2.7 | 3.2846 | 2.8 | 2.8 | 3.16 | 2.9 | 0.53035 | [21 empty] | 2.9638
| 1 | 5 | 3.23 | 2.85 | 3.2612 | 2.9 | 2.7 | 3.08 | 3 | 0.55758 | [21 empty] | 33
| 1 | 6 | 3.01 | 2.73 | 3.2633 | 2.8 | 2.8 | 3.02 | 2.8 | 0.57728
| Mean |  | 3.18 | 2.7183 | 3.2596 | 2.8667 | 2.75 | 3.0883 | 2.8833 | 0.55373
| Median |  | 3.195 | 2.715 | 3.2623 | 2.85 | 2.75 | 3.09 | 2.9 | 0.55217
| Std Dev. |  | 0.11593 | 0.085186 | 0.017138 | 0.08165 | 0.054772 | 0.045789 | 0.075277 | 0.01697
| Rel.Std.Dev. |  | 3.646% | 3.134% | 0.5258% | 2.848% | 1.992% | 1.483% | 2.611% | 3.065%
| PDM3 |  | 7.296% | -8.281% | 9.983% | -3.276% | -7.212% | 4.203% | -2.714% | -81.32%
| Z-Score (Absolute) |  | 1.11 | 0.57 | 1.4 | 0.03 | 0.46 | 0.78 | 0.03 | 8.46
| Analytical results for Ge in OREAS 234b (Certified Value 0.088 ppm) | [30 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | < 0.1 | 1.8 | 0.14063 | < 0.2 | 0.3 | 1.6 | 1.1 | 3.3 | 0.09 | 0.09 | 0.09 | 0.08 | < 10 | 0.6 | < 5 | 0.1 | 0.127 | 0.12 | 0.4 | 0.06 | 0.09 | [8 empty] | 1
| 1 | 2 | 0.2 | 1.8 | 0.1387 | < 0.2 | 0.3 | 1.6 | 1.2 | 2.7 | 0.09 | 0.1 | 0.07 | 0.11 | < 10 | 0.7 | < 5 | 0.09 | 0.1952 | 0.09 | 0.4 | 0.06 | 0.1 | [8 empty] | 35
| 1 | 3 | < 0.1 | 1.7 | 0.1445 | < 0.2 | 0.2 | 1.6 | 1.1 | 3 | 0.09 | 0.1 | 0.09 | 0.07 | < 10 | 0.6 | < 5 | 0.07 | 0.2108 | 0.1 | 0.5 | 0.05 | 0.11 | [8 empty] | 16
| 1 | 4 | < 0.1 | 1.8 | 0.14562 | < 0.2 | 0.2 | 1.6 | 1.3 | 2.9 | 0.08 | 0.1 | 0.06 | 0.08 | < 10 | 0.6 | < 5 | 0.09 | 0.1906 | 0.08 | 0.3 | 0.06 | 0.12 | [8 empty] | 0.087583
| 1 | 5 | < 0.1 | 1.7 | 0.15126 | < 0.2 | 0.3 | 1.6 | 1.1 | 3 | 0.08 | 0.09 | 0.09 | 0.08 | < 10 | 0.6 | < 5 | 0.06 | 0.163 | 0.1 | 0.2 | 0.06 | 0.14 | [8 empty] | 34
| 1 | 6 | < 0.1 | 1.6 | 0.1416 | < 0.2 | 0.2 | 1.7 | 1 | 3 | 0.09 | 0.1 | 0.09 | 0.07 | < 10 | 0.6 | < 5 | 0.22 | 0.13113 | 0.13 | 0.2 | 0.05 | 0.12
| Mean |  | 0.2 | 1.7333 | 0.14372 |  | 0.25 | 1.6167 | 1.1333 | 2.9833 | 0.086667 | 0.096667 | 0.081667 | 0.081667 |  | 0.61667 |  | 0.105 | 0.16962 | 0.10333 | 0.33333 | 0.056667 | 0.11333
| Median |  | 0.2 | 1.75 | 0.14305 |  | 0.25 | 1.6 | 1.1 | 3 | 0.09 | 0.1 | 0.09 | 0.08 |  | 0.6 |  | 0.09 | 0.1768 | 0.1 | 0.35 | 0.06 | 0.115
| Std Dev. |  |  | 0.08165 | 0.0044801 |  | 0.054772 | 0.040825 | 0.10328 | 0.19408 | 0.005164 | 0.005164 | 0.013292 | 0.01472 |  | 0.040825 |  | 0.058224 | 0.03502 | 0.018619 | 0.12111 | 0.005164 | 0.017512
| Rel.Std.Dev. |  |  | 4.711% | 3.117% |  | 21.91% | 2.525% | 9.113% | 6.505% | 5.958% | 5.342% | 16.28% | 18.02% |  | 6.62% |  | 55.45% | 20.65% | 18.02% | 36.33% | 9.113% | 15.45%
| PDM3 |  | 128.4% | 1879% | 64.09% |  | 185.4% | 1746% | 1194% | 3306% | -1.047% | 10.37% | -6.755% | -6.755% |  | 604.1% |  | 19.89% | 93.67% | 17.98% | 280.6% | -35.3% | 29.4%
| Z-Score (Absolute) |  | NA | NA | 1.26 | 0.1 | NA | NA | NA | NA | 0.52 | 0.21 | 0.67 | 0.67 | 152.39 | NA | 74.59 | 0.05 | 2.06 | 0 | NA | 1.45 | 0.31
Results from laboratories 1.01, 1.02, 1.07, 1.10, 1.11, 1.13, 1.22 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 234b (Certified Value 1.64 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.55 | 1.6 | 0.71797 | 1.5 | 1.6 | 1.8 | 1.52 | 1.8 | 1.5 | 1.7 | 1.4 | 1.7 | 1.4 | 1.4 | 1.62 | 2 | 1.8579 | 1.7 | 1.51 | 2.1 | 2.2 | 2 | 1.34 | [6 empty] | 1
| 1 | 2 | 1.81 | 1.52 | 0.7356 | 1.6 | 1.5 | 1.7 | 1.62 | 2 | 1.5 | 1.7 | 1.6 | 1.6 | 1.5 | 1.4 | 1.73 | 1.5 | 1.8839 | 1.6 | 1.5024 | 1.9 | 2.14 | 2 | 1.33 | [6 empty] | 19
| 1 | 3 | 1.76 | 1.55 | 0.65968 | 1.7 | 1.6 | 1.7 | 1.62 | 1.8 | 1.5 | 1.6 | 1.6 | 1.5 | 1.5 | 1.4 | 1.65 | 1.5 | 1.8076 | 1.6 | 1.5221 | 2 | 2.07 | 1.9 | 1.39 | [6 empty] | 16
| 1 | 4 | 1.78 | 1.45 | 0.70434 | 1.6 | 1.6 | 1.6 | 1.65 | 1.8 | 1.5 | 1.6 | 1.5 | 1.5 | 1.4 | 1.4 | 1.73 | 2 | 1.8232 | 2 | 1.485 | 2 | 2.28 | 1.9 | 1.52 | [6 empty] | 1.6435
| 1 | 5 | 1.77 | 1.48 | 0.65525 | 1.8 | 1.6 | 1.6 | 1.53 | 1.9 | 1.5 | 1.7 | 1.6 | 1.5 | 1.4 | 1.3 | 1.7 | 2 | 1.9018 | 1.6 | 1.5288 | 1.8 | 2.27 | 1.9 | 1.43 | [6 empty] | 35
| 1 | 6 | 1.9 | 1.5 | 0.68926 | 0.8 | 1.6 | 1.6 | 1.59 | 1.8 | 1.5 | 1.7 | 1.8 | 1.6 | 1.5 | 1.3 | 1.67 | 1.5 | 1.9103 | 1.5 | 1.4673 | 1.8 | 2.21 | 2.1 | 1.58
| Mean |  | 1.7617 | 1.5167 | 0.69368 | 1.5 | 1.5833 | 1.6667 | 1.5883 | 1.85 | 1.5 | 1.6667 | 1.5833 | 1.5667 | 1.45 | 1.3667 | 1.6833 | 1.75 | 1.8641 | 1.6667 | 1.5026 | 1.9333 | 2.195 | 1.9667 | 1.4317
| Median |  | 1.775 | 1.51 | 0.6968 | 1.6 | 1.6 | 1.65 | 1.605 | 1.8 | 1.5 | 1.7 | 1.6 | 1.55 | 1.45 | 1.4 | 1.685 | 1.75 | 1.8709 | 1.6 | 1.5062 | 1.95 | 2.205 | 1.95 | 1.41
| Std Dev. |  | 0.11548 | 0.053166 | 0.03198 | 0.35777 | 0.040825 | 0.08165 | 0.052694 | 0.083666 | 0 | 0.05164 | 0.13292 | 0.08165 | 0.054772 | 0.05164 | 0.044572 | 0.27386 | 0.042069 | 0.17512 | 0.02314 | 0.12111 | 0.079687 | 0.08165 | 0.10028
| Rel.Std.Dev. |  | 6.555% | 3.505% | 4.61% | 23.85% | 2.578% | 4.899% | 3.318% | 4.522% | 0% | 3.098% | 8.395% | 5.212% | 3.777% | 3.779% | 2.648% | 15.65% | 2.257% | 10.51% | 1.54% | 6.264% | 3.63% | 4.152% | 7.005%
| PDM3 |  | 7.188% | -7.719% | -57.79% | -8.733% | -3.662% | 1.408% | -3.358% | 12.56% | -8.733% | 1.408% | -3.662% | -4.677% | -11.78% | -16.85% | 2.422% | 6.478% | 13.42% | 1.408% | -8.575% | 17.63% | 33.55% | 19.66% | -12.89%
| Z-Score (Absolute) |  | 1.23 | 0.51 | 6.35 | 0.63 | 0.04 | 0.56 | 0 | 1.86 | 0.63 | 0.56 | 0.04 | 0.15 | 0.98 | 1.57 | 0.67 | 1.15 | 1.96 | 0.56 | 0.61 | 2.45 | 4.31 | 2.69 | 1.11
| Analytical results for Hg in OREAS 234b (Indicative Value 0.047 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.05 | 1.25 | 1.30 | [26 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 3
3
| 1 | 1 | < 1 | 0.166 | 0.0198 | [26 empty] | 1
| 1 | 2 | < 1 | 0.2349 | 0.0239 | [26 empty] | 4
| 1 | 3 | < 1 | < 0.1 | 0.0274 | [26 empty] | 16
| 1 | 4 | < 1 | < 0.1 | 0.0316 | [26 empty] | 0.047033
| 1 | 5 | < 1 | < 0.1 | 0.0175 | [26 empty] | 10
| 1 | 6 | < 1 | < 0.1 | < 0.01
| Mean |  |  | 0.20045 | 0.02404
| Median |  |  | 0.20045 | 0.0239
| Std Dev. |  |  | 0.04872 | 0.0056836
| Rel.Std.Dev. |  |  | 24.31% | 23.64%
| PDM3 |  |  | 326.2% | -48.89%
| Z-Score (Absolute) |  | 3.4 | 0 | 0.67
| Analytical results for Ho in OREAS 234b (Certified Value 0.74 ppm) | [30 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.78 | 0.67 | 0.69042 | 0.7 | 0.72 | 0.73 | 0.75 | 0.78205 | [21 empty] | 1
| 1 | 2 | 0.79 | 0.69 | 0.69115 | 0.8 | 0.69 | 0.75 | 0.75 | 0.79905 | [21 empty] | 20
| 1 | 3 | 0.78 | 0.66 | 0.67832 | 0.7 | 0.69 | 0.76 | 0.7 | 0.77097 | [21 empty] | 16
| 1 | 4 | 0.76 | 0.68 | 0.68371 | 0.7 | 0.73 | 0.76 | 0.7 | 0.74423 | [21 empty] | 0.73643
| 1 | 5 | 0.82 | 0.72 | 0.71287 | 0.7 | 0.71 | 0.76 | 0.8 | 0.79591 | [21 empty] | 36
| 1 | 6 | 0.76 | 0.68 | 0.71083 | 0.7 | 0.71 | 0.77 | 0.75 | 0.82065
| Mean |  | 0.78167 | 0.68333 | 0.69455 | 0.71667 | 0.70833 | 0.755 | 0.74167 | 0.78548
| Median |  | 0.78 | 0.68 | 0.69079 | 0.7 | 0.71 | 0.76 | 0.75 | 0.78898
| Std Dev. |  | 0.022286 | 0.020656 | 0.014213 | 0.040825 | 0.016021 | 0.013784 | 0.037639 | 0.026282
| Rel.Std.Dev. |  | 2.851% | 3.023% | 2.046% | 5.696% | 2.262% | 1.826% | 5.075% | 3.346%
| PDM3 |  | 6.142% | -7.21% | -5.687% | -2.684% | -3.816% | 2.521% | 0.7108% | 6.66%
| Z-Score (Absolute) |  | 0.67 | 0.98 | 0.79 | NA | 0.56 | 0.22 | 0 | 0.74
| Analytical results for In in OREAS 234b (Certified Value 0.075 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.08 | 0.08 | 0.066454 | < 0.1 | 0.07 | 0.08 | 0.08 | 0.083 | 0.077 | 0.07 | 0.079 | < 10 | 0.06 | 0.07 | 0.073 | 43.23 | < 0.1 | 0.085 | 0.047 | 0.085 | 0.082 | 0.08 | 0.082 | 0.065 | [5 empty] | 1
| 1 | 2 | 0.08 | 0.08 | 0.060865 | < 0.1 | 0.06 | 0.07 | 0.07 | 0.078 | 0.074 | 0.083 | 0.082 | < 10 | 0.07 | 0.07 | 0.073 | 26.68 | < 0.1 | 0.08509 | 0.038 | 0.0828 | 0.08 | 0.07 | 0.085 | 0.066 | [5 empty] | 21
| 1 | 3 | 0.08 | 0.08 | 0.062853 | < 0.1 | 0.07 | 0.07 | 0.09 | 0.073 | 0.071 | 0.072 | 0.082 | < 10 | 0.06 | 0.07 | 0.073 | 28.45 | < 0.1 | 0.08825 | 0.05 | 0.0794 | 0.073 | 0.08 | 0.077 | 0.069 | [5 empty] | 16
| 1 | 4 | 0.09 | 0.07 | 0.066702 | < 0.1 | 0.07 | 0.07 | 0.09 | 0.08 | 0.071 | 0.072 | 0.077 | < 10 | 0.07 | 0.07 | 0.071 | 24.5 | < 0.1 | 0.08303 | 0.054 | 0.0808 | 0.078 | 0.07 | 0.079 | 0.073 | [5 empty] | 0.074811
| 1 | 5 | 0.08 | 0.07 | 0.055274 | < 0.1 | 0.06 | 0.07 | 0.07 | 0.073 | 0.072 | 0.082 | 0.077 | < 10 | 0.07 | 0.07 | 0.072 | 28.53 | < 0.1 | 0.08268 | 0.043 | 0.0812 | 0.077 | 0.07 | 0.079 | 0.068 | [5 empty] | 37
| 1 | 6 | 0.1 | 0.07 | 0.063597 | < 0.1 | 0.06 | 0.07 | 0.08 | 0.075 | 0.074 | 0.082 | 0.087 | < 10 | 0.07 | 0.07 | 0.076 | 25.18 | < 0.1 | 0.08232 | 0.068 | 0.0791 | 0.074 | 0.08 | 0.08 | 0.075
| Mean |  | 0.085 | 0.075 | 0.062624 |  | 0.065 | 0.071667 | 0.08 | 0.077 | 0.073167 | 0.076833 | 0.080667 |  | 0.066667 | 0.07 | 0.073 | 29.4283 |  | 0.084395 | 0.05 | 0.081383 | 0.077333 | 0.075 | 0.080333 | 0.069333
| Median |  | 0.08 | 0.075 | 0.063225 |  | 0.065 | 0.07 | 0.08 | 0.0765 | 0.073 | 0.077 | 0.0805 |  | 0.07 | 0.07 | 0.073 | 27.565 |  | 0.084015 | 0.0485 | 0.081 | 0.0775 | 0.075 | 0.0795 | 0.0685
| Std Dev. |  | 0.0083666 | 0.0054772 | 0.0042288 |  | 0.0054772 | 0.0040825 | 0.0089443 | 0.0040497 | 0.0023166 | 0.00608 | 0.0038297 |  | 0.005164 | 0 | 0.0016733 | 6.959 |  | 0.002228 | 0.010412 | 0.0022185 | 0.0034448 | 0.0054772 | 0.0028048 | 0.0039328
| Rel.Std.Dev. |  | 9.843% | 7.303% | 6.753% |  | 8.427% | 5.696% | 11.18% | 5.259% | 3.166% | 7.913% | 4.748% |  | 7.746% | 0% | 2.292% | 23.65% |  | 2.64% | 20.82% | 2.726% | 4.454% | 7.303% | 3.491% | 5.672%
| PDM3 |  | 13.62% | 0.2531% | -16.29% |  | -13.11% | -4.203% | 6.937% | 2.927% | -2.198% | 2.704% | 7.828% |  | -10.89% | -6.43% | -2.42% | 39237% |  | 12.81% | -33.16% | 8.786% | 3.372% | 0.2531% | 7.382% | -7.322%
| Z-Score (Absolute) |  | 1.19 | 0 | 1.47 | 2.97 | 1.19 | 0.4 | 0.59 | 0.24 | 0.22 | 0.22 | 0.67 | 586.05 | 0.99 | 0.59 | 0.24 | 3492.92 | 2.97 | 1.12 | 2.97 | 0.76 | 0.28 | 0 | 0.63 | 0.67
| Analytical results for K in OREAS 234b (Certified Value 0.578 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.61 | 0.5749 | 0.56488 | 0.54079 | 0.57 | 0.6284 | 0.57 | 0.5395 | 0.56 | 0.61 | 0.59 | 0.59 | 0.58 | 0.55 | 0.62 | 0.574 | 0.54 | 0.62 | 0.906 | 0.5781 | 0.555 | 0.56 | 0.6412 | 0.59 | 0.64 | 0.6 | 0.58 |  |  | 1
| 1 | 2 | 0.61 | 0.5736 | 0.5637 | 0.5709 | 0.63 | 0.6359 | 0.56 | 0.5589 | 0.55 | 0.63 | 0.58 | 0.58 | 0.54 | 0.55 | 0.61 | 0.583 | 0.54 | 0.6 | 0.852 | 0.57202 | 0.555 | 0.56 | 0.6271 | 0.56 | 0.62 | 0.6 | 0.56 |  |  | #N/A
| 1 | 3 | 0.6 | 0.5705 | 0.56929 | 0.58111 | 0.61 | 0.6193 | 0.569 | 0.5409 | 0.55 | 0.63 | 0.6 | 0.56 | 0.56 | 0.56 | 0.61 | 0.577 | 0.55 | 0.59 | 0.911 | 0.55826 | 0.555 | 0.57 | 0.6413 | 0.57 | 0.63 | 0.61 | 0.56 |  |  | 16
| 1 | 4 | 0.6 | 0.5694 | 0.56345 | 0.56455 | 0.56 | 0.6193 | 0.566 | 0.5603 | 0.57 | 0.61 | 0.58 | 0.56 | 0.53 | 0.55 | 0.57 | 0.575 | 0.54 | 0.6 | 0.969 | 0.57984 | 0.55 | 0.56 | 0.614 | 0.57 | 0.62 | 0.6 | 0.57 |  |  | 0.57775
| 1 | 5 | 0.58 | 0.5622 | 0.56409 | 0.52953 | 0.54 | 0.6201 | 0.572 | 0.5541 | 0.59 | 0.61 | 0.57 | 0.57 | 0.58 | 0.56 | 0.6 | 0.583 | 0.52 | 0.63 | 1.041 | 0.57479 | 0.55 | 0.57 | 0.6386 | 0.58 | 0.63 | 0.6 | 0.61 |  |  | 38
| 1 | 6 | 0.61 | 0.5749 | 0.56779 | 0.56814 | 0.56 | 0.611 | 0.558 | 0.5447 | 0.56 | 0.6 | 0.59 | 0.58 | 0.56 | 0.56 | 0.56 | 0.576 | 0.53 | 0.61 | 0.774 | 0.55998 | 0.555 | 0.54 | 0.636 | 0.57 | 0.61 | 0.61 | 0.58
| Mean |  | 0.60167 | 0.57092 | 0.56553 | 0.55917 | 0.57833 | 0.62233 | 0.56583 | 0.54973 | 0.56333 | 0.615 | 0.585 | 0.57333 | 0.55833 | 0.555 | 0.595 | 0.578 | 0.53667 | 0.60833 | 0.90883 | 0.5705 | 0.55333 | 0.56 | 0.63303 | 0.57333 | 0.625 | 0.60333 | 0.57667
| Median |  | 0.605 | 0.57205 | 0.56448 | 0.56635 | 0.565 | 0.6197 | 0.5675 | 0.5494 | 0.56 | 0.61 | 0.585 | 0.575 | 0.56 | 0.555 | 0.605 | 0.5765 | 0.54 | 0.605 | 0.9085 | 0.57341 | 0.555 | 0.56 | 0.6373 | 0.57 | 0.625 | 0.6 | 0.575
| Std Dev. |  | 0.01169 | 0.0048454 | 0.0024257 | 0.019722 | 0.034303 | 0.0086336 | 0.0056716 | 0.009196 | 0.015055 | 0.012247 | 0.010488 | 0.012111 | 0.020412 | 0.0054772 | 0.02429 | 0.004 | 0.010328 | 0.01472 | 0.092207 | 0.0092324 | 0.002582 | 0.010954 | 0.010697 | 0.010328 | 0.010488 | 0.005164 | 0.018619
| Rel.Std.Dev. |  | 1.943% | 0.8487% | 0.4289% | 3.527% | 5.931% | 1.387% | 1.002% | 1.673% | 2.673% | 1.991% | 1.793% | 2.112% | 3.656% | 0.9869% | 4.082% | 0.692% | 1.924% | 2.42% | 10.15% | 1.618% | 0.4666% | 1.956% | 1.69% | 1.801% | 1.678% | 0.8559% | 3.229%
| PDM3 |  | 4.14% | -1.182% | -2.114% | -3.215% | 0.1014% | 7.717% | -2.062% | -4.849% | -2.495% | 6.448% | 1.255% | -0.764% | -3.36% | -3.937% | 2.986% | 0.04369% | -7.111% | 5.294% | 57.31% | -1.255% | -4.226% | -3.072% | 9.569% | -0.764% | 8.179% | 4.429% | -0.1871%
| Z-Score (Absolute) |  | 1.27 | 0.11 | 0.35 | 0.64 | 0.22 | 2.2 | 0.34 | 1.06 | 0.45 | 1.87 | 0.52 | 0 | 0.67 | 0.82 | 0.97 | 0.21 | 1.65 | 1.57 | 15.08 | 0.13 | 0.9 | 0.6 | 2.68 | 0 | 2.32 | 1.35 | 0.15
| Analytical results for La in OREAS 234b (Certified Value 5.61 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 5.2 | 5.55 | 5.07 | 5.56 | 5.1976 | 5.3 | 6 | 5.34 | 5 | 8 | 5.5 | 6.4 | 5.4 | 5.7 | 6 | 5.1 | 5.4 | 6.5 | 5.5536 | 5.7 | 5.6959 | 6.7 | 5.7 | 6.3 | 5.3 |  |  |  |  | 1
| 1 | 2 | 5.6 | 5.8 | 5.22 | 5.56 | 5.5152 | 5.6 | 6 | 5.35 | 5 | 9 | 5.3 | 6.1 | 5.7 | 5.7 | 6 | 5.1 | 5.3 | 6 | 5.6228 | 5.7 | 5.7287 | 6.7 | 5.5 | 6.5 | 5.3 |  |  |  |  | 22
| 1 | 3 | 6.5 | 5.46 | 4.91 | 5.57 | 5.4847 | 5.6 | 6 | 5.38 | 5 | 9 | 5.4 | 6 | 5.5 | 5.4 | 6 | 5 | 5.2 | 7 | 5.3684 | 5.7 | 5.6836 | 6.7 | 5.3 | 6.1 | 5.5 |  |  |  |  | 16
| 1 | 4 | 6.7 | 5.43 | 5.17 | 5.58 | 5.4649 | 5.2 | 6 | 5.48 | 5 | 9 | 5.2 | 6 | 5.4 | 5.6 | 6 | 5.2 | 5.2 | 6 | 5.3472 | 5.6 | 5.5659 | 6.7 | 5.7 | 6.2 | 6.2 |  |  |  |  | 5.6128
| 1 | 5 | 6 | 5.53 | 5.26 | 5.63 | 5.1996 | 5.2 | 6 | 5.32 | 5 | 9 | 5.3 | 6.2 | 6.6 | 5.5 | 6 | 4.9 | 5.6 | 6 | 5.5465 | 5.2 | 5.5827 | 6.8 | 5.6 | 6.3 | 5.8 |  |  |  |  | 39
| 1 | 6 | 6.1 | 5.33 | 5.2 | 5.62 | 5.4424 | 5.3 | 6 | 5.41 | 6 | 9 | 5.2 | 6 | 6 | 5.7 | 6 | 5 | 5.4 | 6.5 | 5.6345 | 5.5 | 5.4569 | 6.6 | 5.6 | 6.4 | 6.6
| Mean |  | 6.0167 | 5.5167 | 5.1383 | 5.5867 | 5.3841 | 5.3667 | 6 | 5.38 | 5.1667 | 8.8333 | 5.3167 | 6.1167 | 5.7667 | 5.6 | 6 | 5.05 | 5.35 | 6.3333 | 5.5122 | 5.5667 | 5.6189 | 6.7 | 5.5667 | 6.3 | 5.7833
| Median |  | 6.05 | 5.495 | 5.185 | 5.575 | 5.4537 | 5.3 | 6 | 5.365 | 5 | 9 | 5.3 | 6.05 | 5.6 | 5.65 | 6 | 5.05 | 5.35 | 6.25 | 5.5501 | 5.65 | 5.6332 | 6.7 | 5.6 | 6.3 | 5.65
| Std Dev. |  | 0.55648 | 0.15946 | 0.12891 | 0.030768 | 0.14564 | 0.18619 | 0 | 0.05831 | 0.40825 | 0.40825 | 0.1169 | 0.16021 | 0.46762 | 0.12649 | 0 | 0.10488 | 0.15166 | 0.40825 | 0.12488 | 0.19664 | 0.1025 | 0.063246 | 0.15055 | 0.14142 | 0.52694
| Rel.Std.Dev. |  | 9.249% | 2.89% | 2.509% | 0.5507% | 2.705% | 3.469% | 0% | 1.084% | 7.902% | 4.622% | 2.199% | 2.619% | 8.109% | 2.259% | 0% | 2.077% | 2.835% | 6.446% | 2.266% | 3.532% | 1.824% | 0.944% | 2.705% | 2.245% | 9.111%
| PDM3 |  | 7.196% | -1.713% | -8.453% | -0.4655% | -4.075% | -4.385% | 6.899% | -4.148% | -7.948% | 57.38% | -5.276% | 8.977% | 2.741% | -0.2279% | 6.899% | -10.03% | -4.682% | 12.84% | -1.793% | -0.8218% | 0.1097% | 19.37% | -0.8218% | 12.24% | 3.038%
| Z-Score (Absolute) |  | 1.52 | 0.17 | 1.44 | 0.07 | 0.62 | 0.67 | NA | 0.63 | NA | NA | 0.84 | 1.85 | 0.67 | 0.11 | NA | 1.74 | 0.73 | 2.58 | 0.18 | 0 | 0.18 | 3.82 | 0 | 2.47 | 0.73
| Analytical results for Li in OREAS 234b (Certified Value 12.2 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | < 1 | 12.6 | 12.66 | 7.1464 | 12.2121 | 12 | 6 | 13.5 | 11.7 | 12 | 12.3 | 13.8 | 13 | 13.4 | 12.7 | < 5 | 10.5 | 10.8 | 12 | 13 | 10.9091 | 12.5 | 10.7874 | 12.8 | 16 | 12.8 | 11.3 |  |  | 1
| 1 | 2 | < 1 | 12.3 | 12.06 | 7.2474 | 13.0565 | 12 | 6 | 13.3 | 12.2 | 11 | 14.1 | 12.6 | 12.6 | 12.6 | 12.5 | < 5 | 11.8 | 10.8 | 12 | 12 | 11.1236 | 12.5 | 10.6158 | 12.7 | 15 | 12.6 | 10.9 |  |  | #N/A
| 1 | 3 | < 1 | 12.3 | 12.18 | 7.2307 | 13.4453 | 12 | 6 | 13.3 | 12.1 | 11 | 13.4 | 13 | 12.2 | 12.9 | 12.5 | < 5 | 10.5 | 10.8 | 11 | 12 | 10.9381 | 12.6 | 10.7874 | 12.7 | 15 | 12.2 | 11.9 |  |  | 16
| 1 | 4 | < 1 | 12.1 | 11.85 | 7.1535 | 12.9157 | 12 | 6 | 13.7 | 12.2 | 11 | 13.6 | 12.8 | 12.2 | 12.4 | 12.4 | < 5 | 11.1 | 10.8 | 12 | 12 | 10.9854 | 12.3 | 10.413 | 12.6 | 15 | 11.9 | 11.5 |  |  | 12.1613
| 1 | 5 | < 1 | 11.9 | 11.83 | 7.1958 | 12.0825 | 12 | 6 | 13.6 | 12 | 12 | 12.8 | 12.4 | 12.5 | 13.5 | 12.9 | < 5 | 11.3 | 10.6 | 12 | 13 | 11.1452 | 12.4 | 10.7328 | 12.7 | 16 | 12.1 | 11.6 |  |  | 40
| 1 | 6 | < 1 | 11.8 | 11.39 | 7.2098 | 12.9078 | 12 | 6 | 13.2 | 12 | 12 | 15.7 | 12.8 | 12.5 | 13 | 12.8 | < 5 | 11.6 | 10.7 | 12 | 13 | 11.3731 | 12 | 10.7328 | 12.6 | 15 | 12 | 12.6
| Mean |  |  | 12.1667 | 11.995 | 7.1973 | 12.77 | 12 | 6 | 13.4333 | 12.0333 | 11.5 | 13.65 | 12.9 | 12.5 | 12.9667 | 12.6333 |  | 11.1333 | 10.75 | 11.8333 | 12.5 | 11.0791 | 12.3833 | 10.6782 | 12.6833 | 15.3333 | 12.2667 | 11.6333
| Median |  |  | 12.2 | 11.955 | 7.2028 | 12.9118 | 12 | 6 | 13.4 | 12.05 | 11.5 | 13.5 | 12.8 | 12.5 | 12.95 | 12.6 |  | 11.2 | 10.8 | 12 | 12.5 | 11.0545 | 12.45 | 10.7328 | 12.7 | 15 | 12.15 | 11.55
| Std Dev. |  |  | 0.29439 | 0.42307 | 0.040726 | 0.52199 | 0 | 0 | 0.19664 | 0.18619 | 0.54772 | 1.1845 | 0.4858 | 0.29665 | 0.43205 | 0.19664 |  | 0.5465 | 0.083666 | 0.40825 | 0.54772 | 0.17341 | 0.2137 | 0.14425 | 0.075277 | 0.5164 | 0.3559 | 0.5785
| Rel.Std.Dev. |  |  | 2.42% | 3.527% | 0.5658% | 4.088% | 0% | 0% | 1.464% | 1.547% | 4.763% | 8.678% | 3.766% | 2.373% | 3.332% | 1.557% |  | 4.909% | 0.7783% | 3.45% | 4.382% | 1.565% | 1.726% | 1.351% | 0.5935% | 3.368% | 2.901% | 4.973%
| PDM3 |  |  | 0.04374% | -1.368% | -40.82% | 5.005% | -1.327% | -50.66% | 10.46% | -1.053% | -5.438% | 12.24% | 6.074% | 2.785% | 6.622% | 3.881% |  | -8.453% | -11.61% | -2.697% | 2.785% | -8.899% | 1.825% | -12.2% | 4.292% | 26.08% | 0.866% | -4.342%
| Z-Score (Absolute) |  | 10.64 | 0.06 | 0.1 | 4.5 | 0.61 | 0.09 | NA | 1.22 | 0.06 | NA | 1.42 | 0.73 | 0.37 | 0.8 | 0.49 | 8.81 | 0.89 | 1.24 | NA | NA | 0.94 | 0.26 | 1.3 | 0.54 | NA | 0.15 | 0.43
Results from laboratories 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 234b (Certified Value 0.31 ppm) | [30 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | 1.32 | [20 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [20 empty] | 2
3
| 1 | 1 | 0.33 | 0.31 | 0.26 | 4.1983 | 0.3 | 0.33 | 0.5 | 0.33171 | 0.31 | [20 empty] | 1
| 1 | 2 | 0.35 | 0.33 | 0.28 | 4.1653 | 0.3 | 0.33 | 0.4 | 0.33788 | 0.29 | [20 empty] | 23
| 1 | 3 | 0.35 | 0.32 | 0.27 | 4.1898 | 0.3 | 0.32 | 0.4 | 0.31442 | 0.27 | [20 empty] | 16
| 1 | 4 | 0.35 | 0.3 | 0.27 | 4.2222 | 0.3 | 0.33 | 0.5 | 0.32187 | 0.3 | [20 empty] | 0.31269
| 1 | 5 | 0.33 | 0.32 | 0.28 | 4.2053 | 0.3 | 0.31 | 0.4 | 0.34337 | 0.29 | [20 empty] | 41
| 1 | 6 | 0.35 | 0.3 | 0.28 | 4.1469 | 0.3 | 0.33 | 0.5 | 0.32763 | 0.29
| Mean |  | 0.34333 | 0.31333 | 0.27333 | 4.188 | 0.3 | 0.325 | 0.45 | 0.32948 | 0.29167
| Median |  | 0.35 | 0.315 | 0.275 | 4.1941 | 0.3 | 0.33 | 0.45 | 0.32967 | 0.29
| Std Dev. |  | 0.010328 | 0.012111 | 0.008165 | 0.0275 | 0 | 0.0083666 | 0.054772 | 0.010551 | 0.013292
| Rel.Std.Dev. |  | 3.008% | 3.865% | 2.987% | 0.6566% | 0% | 2.574% | 12.17% | 3.202% | 4.557%
| PDM3 |  | 9.8% | 0.2054% | -12.59% | 1239% | -4.059% | 3.936% | 43.91% | 5.369% | -6.724%
| Z-Score (Absolute) |  | 0.67 | 0.43 | 1.9 | 142.08 | NA | 0 | NA | 0.16 | 1.23
Results from laboratories 1.07 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 234b (Certified Value 3.76 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.7 | 3.7232 | 3.8901 | 3.5102 | 3.56 | 4.137 | 3.89 | 3.6074 | 3.49 | 3.28 | 4.02 | 3.93 | 3.96 | 3.77 | 3.94 | 3.847 | 3.88 | 3.781 | 1.68 | 3.2058 | 3.934 | 3.8 | 3.907 | 3.92 | 3.08 | 3.82 | 3.7 |  |  | 1
| 1 | 2 | 3.7 | 3.7177 | 3.9136 | 3.7095 | 3.73 | 4.1672 | 3.82 | 3.7862 | 3.45 | 3.42 | 3.86 | 3.83 | 3.69 | 3.76 | 3.89 | 3.829 | 3.83 | 3.625 | 1.06 | 3.1077 | 3.9455 | 3.77 | 3.8449 | 3.78 | 2.92 | 3.8 | 3.6 |  |  | #N/A
| 1 | 3 | 3.67 | 3.6309 | 3.8927 | 3.7862 | 3.75 | 4.0767 | 3.89 | 3.7354 | 3.43 | 3.44 | 3.97 | 3.7 | 3.78 | 3.78 | 3.88 | 3.8 | 3.85 | 3.619 | 1.18 | 3.0944 | 3.9465 | 3.85 | 3.8967 | 3.8 | 2.97 | 3.84 | 3.7 |  |  | 16
| 1 | 4 | 3.7 | 3.7556 | 3.892 | 3.6829 | 3.46 | 4.0586 | 3.82 | 3.7803 | 3.41 | 3.44 | 3.87 | 3.68 | 3.64 | 3.74 | 3.67 | 3.862 | 3.87 | 3.622 | 0.882 | 3.1374 | 3.9085 | 3.79 | 3.7618 | 3.8 | 2.95 | 3.82 | 3.7 |  |  | 3.7644
| 1 | 5 | 3.59 | 3.6264 | 3.936 | 3.4456 | 3.47 | 4.1129 | 3.89 | 3.7705 | 3.47 | 3.4 | 3.81 | 3.73 | 3.91 | 3.84 | 3.89 | 3.839 | 3.79 | 3.831 | 1.28 | 3.1652 | 3.934 | 3.85 | 3.8919 | 3.86 | 3.01 | 3.78 | 3.4 |  |  | 42
| 1 | 6 | 3.76 | 3.554 | 3.9036 | 3.7001 | 3.49 | 4.0586 | 3.83 | 3.7187 | 3.49 | 3.37 | 3.86 | 3.78 | 3.77 | 3.82 | 3.61 | 3.807 | 3.83 | 3.698 | 0.9 | 3.0678 | 3.9285 | 3.7 | 3.9 | 3.79 | 2.95 | 3.81 | 4.3
| Mean |  | 3.6867 | 3.668 | 3.9047 | 3.6391 | 3.5767 | 4.1018 | 3.8567 | 3.7331 | 3.4567 | 3.3917 | 3.8983 | 3.775 | 3.7917 | 3.785 | 3.8133 | 3.8307 | 3.8417 | 3.696 | 1.1637 | 3.1297 | 3.9328 | 3.7933 | 3.867 | 3.825 | 2.98 | 3.8117 | 3.7333
| Median |  | 3.7 | 3.6743 | 3.8982 | 3.6915 | 3.525 | 4.0948 | 3.86 | 3.753 | 3.46 | 3.41 | 3.865 | 3.755 | 3.775 | 3.775 | 3.885 | 3.834 | 3.84 | 3.6615 | 1.12 | 3.1225 | 3.934 | 3.795 | 3.8943 | 3.8 | 2.96 | 3.815 | 3.7
| Std Dev. |  | 0.055737 | 0.076543 | 0.017757 | 0.13135 | 0.1314 | 0.044726 | 0.036697 | 0.067021 | 0.03266 | 0.0608 | 0.079352 | 0.093541 | 0.12352 | 0.037815 | 0.13721 | 0.023754 | 0.032506 | 0.091542 | 0.29665 | 0.050347 | 0.013862 | 0.056095 | 0.056121 | 0.054314 | 0.057271 | 0.020412 | 0.30111
| Rel.Std.Dev. |  | 1.512% | 2.087% | 0.4548% | 3.609% | 3.674% | 1.09% | 0.9515% | 1.795% | 0.9448% | 1.793% | 2.036% | 2.478% | 3.258% | 0.9991% | 3.598% | 0.6201% | 0.8462% | 2.477% | 25.49% | 1.609% | 0.3525% | 1.479% | 1.451% | 1.42% | 1.922% | 0.5355% | 8.065%
| PDM3 |  | -2.065% | -2.562% | 3.725% | -3.329% | -4.988% | 8.963% | 2.451% | -0.8324% | -8.175% | -9.902% | 3.557% | 0.2811% | 0.7239% | 0.5468% | 1.299% | 1.76% | 2.052% | -1.817% | -69.09% | -16.86% | 4.474% | 0.7681% | 2.726% | 1.609% | -20.84% | 1.255% | -0.8257%
| Z-Score (Absolute) |  | 0.75 | 0.89 | 0.91 | 1.11 | 1.58 | 2.4 | 0.54 | 0.39 | 2.49 | 2.98 | 0.86 | 0.08 | 0.05 | 0 | 0.21 | 0.35 | 0.43 | 0.67 | 19.86 | 4.96 | 1.12 | 0.06 | 0.62 | 0.3 | 6.1 | 0.2 | 0.39
| Analytical results for Mn in OREAS 234b (Certified Value 0.148 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.15 | 0.1512 | 0.15499 | 0.13783 | 0.145 | 0.1688 | 0.149 | 0.1443 | 0.1345 | 0.1648 | 0.151 | 0.1525 | 0.1575 | 0.147 | 0.1522 | 0.1496 | 0.1524 | 0.1512 | 0.116 | 0.14261 | 0.1455 | 0.142 | 0.13367 | 0.154 | 0.1536 | 0.1451 | 0.144 |  |  | 1
| 1 | 2 | 0.15 | 0.1522 | 0.15457 | 0.14596 | 0.154 | 0.1696 | 0.146 | 0.1515 | 0.1347 | 0.1646 | 0.1445 | 0.15 | 0.1475 | 0.1465 | 0.1508 | 0.1515 | 0.1511 | 0.1454 | 0.081 | 0.13933 | 0.146 | 0.141 | 0.13226 | 0.149 | 0.1487 | 0.1439 | 0.14 |  |  | 24
| 1 | 3 | 0.15 | 0.1506 | 0.15582 | 0.15011 | 0.153 | 0.1657 | 0.149 | 0.148 | 0.1353 | 0.1657 | 0.1495 | 0.1445 | 0.151 | 0.147 | 0.1506 | 0.1506 | 0.1513 | 0.1435 | 0.083 | 0.1364 | 0.146 | 0.143 | 0.13453 | 0.151 | 0.1484 | 0.1459 | 0.142 |  |  | 16
| 1 | 4 | 0.15 | 0.1532 | 0.15474 | 0.1449 | 0.142 | 0.1657 | 0.147 | 0.1518 | 0.135 | 0.1598 | 0.145 | 0.145 | 0.145 | 0.147 | 0.1426 | 0.15 | 0.1518 | 0.1447 | 0.078 | 0.13966 | 0.145 | 0.142 | 0.13021 | 0.151 | 0.1525 | 0.1456 | 0.151 |  |  | 0.14785
| 1 | 5 | 0.15 | 0.1475 | 0.15688 | 0.13605 | 0.141 | 0.1665 | 0.149 | 0.1495 | 0.1349 | 0.1485 | 0.1425 | 0.1475 | 0.1545 | 0.1495 | 0.1503 | 0.1515 | 0.149 | 0.1521 | 0.084 | 0.13945 | 0.1465 | 0.145 | 0.13367 | 0.154 | 0.1516 | 0.144 | 0.156 |  |  | 43
| 1 | 6 | 0.15 | 0.1459 | 0.15863 | 0.14488 | 0.137 | 0.165 | 0.146 | 0.1473 | 0.1353 | 0.1625 | 0.146 | 0.1495 | 0.1515 | 0.1485 | 0.1403 | 0.1515 | 0.1503 | 0.1476 | 0.078 | 0.13561 | 0.146 | 0.1405 | 0.13464 | 0.151 | 0.1494 | 0.1453 | 0.15
| Mean |  | 0.15 | 0.1501 | 0.15594 | 0.14329 | 0.14533 | 0.16688 | 0.14767 | 0.14873 | 0.13495 | 0.16098 | 0.14642 | 0.14817 | 0.15117 | 0.14758 | 0.1478 | 0.15078 | 0.15098 | 0.14742 | 0.086667 | 0.13884 | 0.14583 | 0.14225 | 0.13316 | 0.15167 | 0.1507 | 0.14497 | 0.14717
| Median |  | 0.15 | 0.1509 | 0.1554 | 0.14489 | 0.1435 | 0.1661 | 0.148 | 0.14875 | 0.13495 | 0.16355 | 0.1455 | 0.1485 | 0.15125 | 0.147 | 0.15045 | 0.15105 | 0.1512 | 0.1465 | 0.082 | 0.13939 | 0.146 | 0.142 | 0.13367 | 0.151 | 0.1505 | 0.1452 | 0.147
| Std Dev. |  | 0 | 0.0028242 | 0.0015734 | 0.0053106 | 0.0068313 | 0.0018734 | 0.0015055 | 0.0028247 | 0.00032094 | 0.0064719 | 0.0032158 | 0.003093 | 0.004535 | 0.0011583 | 0.0050148 | 0.00084715 | 0.001199 | 0.0035516 | 0.014583 | 0.0025272 | 0.0005164 | 0.0016047 | 0.0016774 | 0.0019664 | 0.0021652 | 0.00083347 | 0.0061455
| Rel.Std.Dev. |  | 0% | 1.882% | 1.009% | 3.706% | 4.7% | 1.123% | 1.02% | 1.899% | 0.2378% | 4.02% | 2.196% | 2.088% | 3% | 0.7848% | 3.393% | 0.5618% | 0.7941% | 2.409% | 16.83% | 1.82% | 0.3541% | 1.128% | 1.26% | 1.297% | 1.437% | 0.5749% | 4.176%
| PDM3 |  | 1.451% | 1.519% | 5.467% | -3.089% | -1.705% | 12.87% | -0.1272% | 0.5943% | -8.728% | 8.879% | -0.9726% | 0.211% | 2.24% | -0.1835% | -0.03699% | 1.981% | 2.116% | -0.2963% | -41.38% | -6.094% | -1.367% | -3.791% | -9.936% | 2.578% | 1.924% | -1.953% | -0.4653%
| Z-Score (Absolute) |  | 0.52 | 0.54 | 1.84 | 0.97 | 0.52 | 4.27 | 0 | 0.24 | 2.83 | 2.96 | 0.28 | 0.11 | 0.78 | 0.02 | 0.03 | 0.69 | 0.74 | 0.06 | 13.56 | 1.96 | 0.41 | 1.2 | 3.22 | 0.89 | 0.67 | 0.6 | 0.11
| Analytical results for Mo in OREAS 234b (Certified Value 2.57 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.86 | 2.6 | 3.1 | 2.4215 | 2.3066 | 2 | 2.92 | 2.6 | 2.4 | 2 | 2.9 | 2.64 | 2.71 | 2.47 | 2.52 | 2 | 2.1 | 2.2 | 2.35 | 3.08 | 3 | 2.5527 | 2.57 | 2.4302 | 2.77 | 2.96 | 2.68 | 2.28 |  | 1
| 1 | 2 | 3.01 | 2.7 | 2.48 | 2.4766 | 2.4301 | 2 | 2.85 | 2.7 | 2.6 | 2 | 2.8 | 2.45 | 2.61 | 2.31 | 2.47 | < 2 | 2.1 | 2.2 | 2.25 | 2.92 | 3 | 2.4338 | 2.5 | 2.4488 | 2.62 | 3.25 | 2.7 | 2.21 |  | 25
| 1 | 3 | 2.84 | 2.7 | 2.32 | 2.4499 | 2.3936 | 3 | 2.88 | 2.6 | 2.6 | 2 | 2.7 | 2.53 | 2.72 | 2.43 | 2.49 | 2 | 2 | 2.2 | 2.48 | 2.96 | 3 | 2.6342 | 2.63 | 2.3916 | 2.72 | 2.84 | 2.53 | 2.26 |  | 16
| 1 | 4 | 2.82 | 2.5 | 2.53 | 2.454 | 2.4019 | 2 | 2.96 | 2.5 | 2.5 | 2 | 2.7 | 2.43 | 2.62 | 2.41 | 2.45 | 2 | 2 | 2.2 | 2.34 | 2.88 | 3 | 2.632 | 2.6 | 2.4002 | 2.75 | 2.8 | 2.62 | 2.45 |  | 2.5666
| 1 | 5 | 2.77 | 2.6 | 2.86 | 2.4526 | 2.235 | 2 | 3 | 2.6 | 2.7 | 2 | 2.1 | 2.43 | 2.64 | 2.73 | 2.57 | 2 | 2 | 2.1 | 2.36 | 2.67 | 3 | 2.5301 | 2.59 | 2.4348 | 2.73 | 2.91 | 2.61 | 2.36 |  | 44
| 1 | 6 | 2.82 | 2.5 | 2.96 | 2.4958 | 2.3341 | < 1 | 2.86 | 2.5 | 2.5 | 3 | 2.8 | 2.41 | 2.78 | 2.6 | 2.49 | < 2 | 2.1 | 2.2 | 2.54 | 2.55 | 3 | 2.5321 | 2.58 | 2.4177 | 2.76 | 2.97 | 2.68 | 2.61
| Mean |  | 2.8533 | 2.6 | 2.7083 | 2.4584 | 2.3502 | 2.2 | 2.9117 | 2.5833 | 2.55 | 2.1667 | 2.6667 | 2.4817 | 2.68 | 2.4917 | 2.4983 | 2 | 2.05 | 2.1833 | 2.3867 | 2.8433 | 3 | 2.5525 | 2.5783 | 2.4205 | 2.725 | 2.955 | 2.6367 | 2.3617
| Median |  | 2.83 | 2.6 | 2.695 | 2.4533 | 2.3639 | 2 | 2.9 | 2.6 | 2.55 | 2 | 2.75 | 2.44 | 2.675 | 2.45 | 2.49 | 2 | 2.05 | 2.2 | 2.355 | 2.9 | 3 | 2.5424 | 2.585 | 2.424 | 2.74 | 2.935 | 2.65 | 2.32
| Std Dev. |  | 0.082381 | 0.089443 | 0.30805 | 0.025385 | 0.072627 | 0.44721 | 0.05947 | 0.075277 | 0.10488 | 0.40825 | 0.28752 | 0.088185 | 0.067231 | 0.14999 | 0.042151 | 0 | 0.054772 | 0.040825 | 0.10501 | 0.19643 | 0 | 0.074824 | 0.043551 | 0.021701 | 0.054681 | 0.15909 | 0.063456 | 0.14798
| Rel.Std.Dev. |  | 2.887% | 3.44% | 11.37% | 1.033% | 3.09% | 20.33% | 2.042% | 2.914% | 4.113% | 18.84% | 10.78% | 3.553% | 2.509% | 6.02% | 1.687% | 0% | 2.672% | 1.87% | 4.4% | 6.909% | 0% | 2.931% | 1.689% | 0.8965% | 2.007% | 5.384% | 2.407% | 6.266%
| PDM3 |  | 11.17% | 1.302% | 5.523% | -4.215% | -8.43% | -14.28% | 13.45% | 0.6526% | -0.6461% | -15.58% | 3.9% | -3.309% | 4.419% | -2.919% | -2.659% | -22.08% | -20.13% | -14.93% | -7.01% | 10.78% | 16.89% | -0.5491% | 0.4578% | -5.69% | 6.172% | 15.13% | 2.731% | -7.984%
| Z-Score (Absolute) |  | 1.51 | 0.18 | 0.75 | 0.56 | 1.13 | NA | 1.81 | 0.09 | 0.08 | NA | 0.53 | 0.44 | 0.6 | 0.39 | 0.35 | NA | 2.7 | 2 | 0.94 | 1.46 | NA | 0.07 | 0.07 | 0.76 | 0.84 | 2.04 | 0.37 | 1.07
Results from laboratories 1.07, 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 234b (Certified Value 1.94 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.99 | 1.9388 | 1.9413 | 1.8187 | 1.8524 | 1.95 | 2.166 | 1.95 | 1.7886 | 1.82 | 1.99 | 2.04 | 2.02 | 2.04 | 1.9 | 2.08 | 1.897 | 1.896 | 1.911 | 2.404 | 1.9752 | 1.99 | 1.97 | 2.581 | 2.01 | 2.41 | 2.08 | 2.01 |  | 1
| 1 | 2 | 2.01 | 1.9818 | 1.833 | 1.8185 | 1.9492 | 2.02 | 2.166 | 1.91 | 1.8793 | 1.8 | 2.06 | 1.99 | 1.95 | 1.9 | 1.9 | 2.06 | 1.972 | 1.946 | 1.844 | 2.294 | 1.9151 | 1.995 | 1.95 | 2.527 | 1.95 | 2.27 | 2.07 | 1.96 |  | #N/A
| 1 | 3 | 1.97 | 1.9314 | 1.9014 | 1.8237 | 1.9829 | 2.01 | 2.129 | 1.95 | 1.8485 | 1.79 | 2.06 | 2.04 | 1.91 | 1.96 | 1.9 | 2.08 | 1.927 | 1.913 | 1.838 | 2.564 | 1.871 | 1.99 | 1.96 | 2.54 | 1.95 | 2.32 | 2.1 | 1.96 |  | 16
| 1 | 4 | 1.97 | 1.9519 | 2.0568 | 1.8104 | 1.9194 | 1.87 | 2.144 | 1.92 | 1.9023 | 1.84 | 2.01 | 1.98 | 1.9 | 1.89 | 1.89 | 1.97 | 1.931 | 1.904 | 1.843 | 2.632 | 1.9321 | 1.985 | 1.96 | 2.428 | 1.97 | 2.31 | 2.08 | 1.86 |  | 1.942
| 1 | 5 | 1.94 | 1.9368 | 2.0318 | 1.8359 | 1.7991 | 1.87 | 2.136 | 1.95 | 1.8685 | 1.81 | 2 | 1.97 | 1.91 | 2.01 | 1.94 | 2.06 | 1.978 | 1.929 | 1.939 | 2.84 | 1.9258 | 1.99 | 1.98 | 2.526 | 2 | 2.39 | 2.06 | 1.99 |  | 45
| 1 | 6 | 2.02 | 1.8798 | 2.0033 | 1.8589 | 1.9464 | 1.89 | 2.136 | 1.89 | 1.8525 | 1.8 | 1.98 | 1.99 | 1.96 | 1.97 | 1.92 | 1.91 | 1.883 | 1.891 | 1.886 | 2.132 | 1.8822 | 1.985 | 1.87 | 2.483 | 1.95 | 2.34 | 2.08 | 1.85
| Mean |  | 1.9833 | 1.9367 | 1.9613 | 1.8277 | 1.9082 | 1.935 | 2.1462 | 1.9283 | 1.8566 | 1.81 | 2.0167 | 2.0017 | 1.9417 | 1.9617 | 1.9083 | 2.0267 | 1.9313 | 1.9132 | 1.8768 | 2.4777 | 1.9169 | 1.9892 | 1.9483 | 2.5142 | 1.9717 | 2.34 | 2.0783 | 1.9383
| Median |  | 1.98 | 1.9378 | 1.9723 | 1.8212 | 1.9329 | 1.92 | 2.14 | 1.935 | 1.8605 | 1.805 | 2.005 | 1.99 | 1.93 | 1.965 | 1.9 | 2.06 | 1.929 | 1.9085 | 1.865 | 2.484 | 1.9204 | 1.99 | 1.96 | 2.5265 | 1.96 | 2.33 | 2.08 | 1.96
| Std Dev. |  | 0.029439 | 0.033272 | 0.085232 | 0.017445 | 0.069064 | 0.068629 | 0.01608 | 0.025626 | 0.038605 | 0.017889 | 0.035024 | 0.030605 | 0.04535 | 0.059133 | 0.018348 | 0.070333 | 0.038381 | 0.020971 | 0.042064 | 0.25346 | 0.037462 | 0.0037639 | 0.039707 | 0.052602 | 0.027142 | 0.052154 | 0.013292 | 0.067355
| Rel.Std.Dev. |  | 1.484% | 1.718% | 4.346% | 0.9545% | 3.619% | 3.547% | 0.7492% | 1.329% | 2.079% | 0.9883% | 1.737% | 1.529% | 2.336% | 3.014% | 0.9615% | 3.47% | 1.987% | 1.096% | 2.241% | 10.23% | 1.954% | 0.1892% | 2.038% | 2.092% | 1.377% | 2.229% | 0.6395% | 3.475%
| PDM3 |  | 2.131% | -0.2681% | 0.9945% | -5.884% | -1.737% | -0.3582% | 10.52% | -0.7015% | -4.394% | -6.795% | 3.847% | 3.075% | -0.01489% | 1.015% | -1.731% | 4.362% | -0.547% | -1.482% | -3.353% | 27.59% | -1.29% | 2.431% | 0.3284% | 29.47% | 1.53% | 20.5% | 7.023% | -0.1865%
| Z-Score (Absolute) |  | 0.69 | 0.15 | 0.29 | 2.11 | 0.66 | 0.18 | 3.62 | 0.3 | 1.59 | 2.42 | 1.29 | 1.02 | 0.06 | 0.3 | 0.66 | 1.47 | 0.25 | 0.57 | 1.22 | 9.56 | 0.5 | 0.79 | 0.06 | 10.22 | 0.48 | 7.09 | 2.39 | 0.12
| Analytical results for Nb in OREAS 234b (Certified Value 3.59 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.5 | 3.78 | 20.78 | 2.2499 | 3.5339 | 1.5 | 4.1 | 3.5 | 3.33 | < 5 | 3.7 | 3.7 | 3.7 | 3.4 | 3.4 | 24 | 3 | 3 | 3.4 | 5 | 3.1954 | 3.7 | 5.2081 | 4.4 | 3.9 | 4 | 3.2 |  |  | 1
| 1 | 2 | 3.7 | 3.72 | 19.39 | 2.2673 | 3.6773 | 3.2 | 3.9 | 3.6 | 3.44 | < 5 | 3.8 | 3.7 | 3.6 | 3.6 | 3.5 | 25 | 3.3 | 3 | 3.4 | 5 | 3.232 | 3.7 | 5.1883 | 4 | 3.6 | 4 | 3.1 |  |  | 8
| 1 | 3 | 3.7 | 3.77 | 20.56 | 2.1963 | 3.7851 | 3.7 | 3.9 | 3.8 | 3.5 | < 5 | 3.7 | 3.8 | 3.6 | 3.5 | 3.5 | 25 | 3.2 | 3 | 3.7 | 5 | 2.9829 | 3.7 | 5.2393 | 4.2 | 3.8 | 3.8 | 3.3 |  |  | 16
| 1 | 4 | 3.6 | 3.72 | 23.11 | 2.1535 | 3.6117 | 3.4 | 4 | 3.6 | 3.65 | < 5 | 3.8 | 3.7 | 3.7 | 3.4 | 3.4 | 23 | 3 | 3.1 | 3.5 | 5 | 2.6499 | 3.7 | 5.0965 | 4.2 | 3.8 | 3.9 | 3.4 |  |  | 3.5911
| 1 | 5 | 3.6 | 3.68 | 20.9 | 2.1993 | 3.3561 | 3.4 | 4 | 3.7 | 3.57 | < 5 | 3.7 | 3.7 | 3.7 | 3.6 | 3.4 | 24 | 3.2 | 3 | 3.4 | 5 | 3.268 | 3.8 | 5.1543 | 3.9 | 3.7 | 3.9 | 3.3 |  |  | 46
| 1 | 6 | 3.6 | 3.58 | 18.39 | 2.2211 | 3.7074 | 0.3 | 4 | 3.6 | 3.48 | < 5 | 3.7 | 3.7 | 3.7 | 4 | 3.5 | 23 | 3.3 | 3 | 3.6 | 5 | 3.3406 | 3.3 | 5.0674 | 4.1 | 3.8 | 4.2 | 3.6
| Mean |  | 3.6167 | 3.7083 | 20.5217 | 2.2146 | 3.6119 | 2.5833 | 3.9833 | 3.6333 | 3.495 |  | 3.7333 | 3.7167 | 3.6667 | 3.5833 | 3.45 | 24 | 3.1667 | 3.0167 | 3.5 | 5 | 3.1115 | 3.65 | 5.159 | 4.1333 | 3.7667 | 3.9667 | 3.3167
| Median |  | 3.6 | 3.72 | 20.67 | 2.2102 | 3.6445 | 3.3 | 4 | 3.6 | 3.49 |  | 3.7 | 3.7 | 3.7 | 3.55 | 3.45 | 24 | 3.2 | 3 | 3.45 | 5 | 3.2137 | 3.7 | 5.1713 | 4.15 | 3.8 | 3.95 | 3.3
| Std Dev. |  | 0.075277 | 0.072778 | 1.596 | 0.040938 | 0.15159 | 1.3674 | 0.075277 | 0.10328 | 0.10968 |  | 0.05164 | 0.040825 | 0.05164 | 0.22286 | 0.054772 | 0.89443 | 0.13663 | 0.040825 | 0.12649 | 0 | 0.25619 | 0.17607 | 0.066387 | 0.17512 | 0.10328 | 0.13663 | 0.17224
| Rel.Std.Dev. |  | 2.081% | 1.963% | 7.777% | 1.849% | 4.197% | 52.93% | 1.89% | 2.843% | 3.138% |  | 1.383% | 1.098% | 1.408% | 6.219% | 1.588% | 3.727% | 4.315% | 1.353% | 3.614% | 0% | 8.234% | 4.824% | 1.287% | 4.237% | 2.742% | 3.444% | 5.193%
| PDM3 |  | 0.7107% | 3.263% | 471.5% | -38.33% | 0.5779% | -28.06% | 10.92% | 1.175% | -2.677% |  | 3.959% | 3.495% | 2.103% | -0.2175% | -3.93% | 568.3% | -11.82% | -16% | -2.538% | 39.23% | -13.36% | 1.639% | 43.66% | 15.1% | 4.888% | 10.46% | -7.643%
| Z-Score (Absolute) |  | 0.02 | 0.23 | 47.15 | 3.94 | 0.04 | 2.91 | 1 | 0.02 | 0.36 | 3.14 | 0.3 | 0.26 | 0.12 | 0.12 | 0.49 | 56.85 | 1.28 | 1.7 | 0.35 | NA | 1.43 | 0.07 | 4.28 | 1.42 | 0.4 | 0.95 | 0.86
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 234b (Certified Value 8.39 ppm) | [30 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 8.79 | 4.8 | 8.014 | 7.9 | 8.9 | 8.2 | 8.5 | 8.7705 | [21 empty] | 1
| 1 | 2 | 8.87 | 5.09 | 7.9588 | 8.4 | 8.6 | 8.2 | 9 | 8.6968 | [21 empty] | 27
| 1 | 3 | 8.66 | 4.69 | 7.9776 | 8.5 | 8.5 | 8 | 8.5 | 8.5308 | [21 empty] | 16
| 1 | 4 | 8.49 | 4.93 | 7.9926 | 8.1 | 8.9 | 8.1 | 8 | 8.4742 | [21 empty] | 8.3934
| 1 | 5 | 8.45 | 5.23 | 7.9506 | 8 | 8.8 | 8 | 8.5 | 8.739 | [21 empty] | 47
| 1 | 6 | 8.2 | 4.88 | 7.904 | 7.9 | 8.7 | 8.5 | 8.5 | 8.8541
| Mean |  | 8.5767 | 4.9367 | 7.9662 | 8.1333 | 8.7333 | 8.1667 | 8.5 | 8.6776
| Median |  | 8.575 | 4.905 | 7.9682 | 8.05 | 8.75 | 8.15 | 8.5 | 8.7179
| Std Dev. |  | 0.24655 | 0.19613 | 0.038146 | 0.2582 | 0.1633 | 0.18619 | 0.31623 | 0.14619
| Rel.Std.Dev. |  | 2.875% | 3.973% | 0.4788% | 3.175% | 1.87% | 2.28% | 3.72% | 1.685%
| PDM3 |  | 2.183% | -41.18% | -5.089% | -3.098% | 4.05% | -2.701% | 1.27% | 3.385%
| Z-Score (Absolute) |  | 0.56 | 7.8 | 0.84 | 0.46 | 0.92 | 0.38 | 0.38 | 0.79
| Analytical results for Ni in OREAS 234b (Certified Value 97 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 97.5 | 100.7 | 94.9 | 96.95 | 97.1083 | 93 | 116 | 99 | 94.2 | 94 | 92.9 | 105 | 100 | 99 | 98.7 | 90 | 97 | 97 | 96 | 97.99 | 94.4787 | 95 | 97.796 | 97.2 | 90.5051 | 89.7 | 106.5 | 101.7 | 101 | 1
| 1 | 2 | 96.5 | 100.3 | 92.17 | 97.28 | 102.8727 | 95 | 112 | 96 | 99.1 | 94 | 93.6 | 103 | 101 | 97.3 | 98.5 | 90 | 94 | 96.2 | 92 | 80.54 | 94.5159 | 95 | 98.9853 | 98.3 | 86.6714 | 89.5 | 100.9 | 100.3 | 98 | 28
| 1 | 3 | 96.8 | 99.8 | 91.2 | 97.84 | 105.1127 | 95 | 113 | 98 | 97.4 | 95 | 96.7 | 106 | 100 | 95.9 | 95.3 | 90 | 95 | 97.1 | 92 | 80.47 | 89.7264 | 95 | 99.1244 | 101 | 86.9833 | 89 | 101.1 | 101.1 | 98.9 | 16
| 1 | 4 | 97.7 | 99.8 | 92.79 | 96.94 | 102.5095 | 93 | 113 | 96 | 97.8 | 99 | 95.6 | 104 | 99.4 | 93.6 | 95.9 | 85 | 95 | 98.7 | 92 | 78.81 | 92.7024 | 95 | 102.6606 | 97.8 | 82.9671 | 90.4 | 105.9 | 101.5 | 97.9 | 96.9143
| 1 | 5 | 94.2 | 96.8 | 92.01 | 97.52 | 95.088 | 95 | 115 | 98 | 98.1 | 93 | 95.9 | 104 | 98.7 | 102.5 | 97.6 | 90 | 96 | 94.4 | 97 | 80.18 | 93.1674 | 95 | 101.0672 | 99.2 | 82.5069 | 88.6 | 101.9 | 100.6 | 101 | 48
| 1 | 6 | 98.3 | 96.2 | 91.99 | 99.37 | 102.1275 | 95 | 115 | 97 | 97.6 | 99 | 94.1 | 105 | 99.8 | 101 | 99.9 | 83 | 94 | 95.6 | 94 | 78.02 | 90.1263 | 95 | 101.3645 | 95.6 | 84.3341 | 90.9 | 103.3 | 100.9 | 96.8
| Mean |  | 96.8333 | 98.9333 | 92.51 | 97.65 | 100.8031 | 94.3333 | 114 | 97.3333 | 97.3667 | 95.6667 | 94.8 | 104.5 | 99.8167 | 98.2167 | 97.65 | 88 | 95.1667 | 96.5 | 93.8333 | 82.6683 | 92.4529 | 95 | 100.1663 | 98.1833 | 85.6613 | 89.6833 | 103.2667 | 101.0167 | 98.9333
| Median |  | 97.15 | 99.8 | 92.09 | 97.4 | 102.3185 | 95 | 114 | 97.5 | 97.7 | 94.5 | 94.85 | 104.5 | 99.9 | 98.15 | 98.05 | 90 | 95 | 96.6 | 93 | 80.325 | 92.9349 | 95 | 100.0958 | 98.05 | 85.5028 | 89.6 | 102.6 | 101 | 98.45
| Std Dev. |  | 1.4418 | 1.9242 | 1.2761 | 0.91021 | 3.8428 | 1.0328 | 1.5492 | 1.2111 | 1.6621 | 2.6583 | 1.4832 | 1.0488 | 0.76004 | 3.2933 | 1.7593 | 3.1623 | 1.169 | 1.467 | 2.2286 | 7.5737 | 2.0871 | 0 | 1.8203 | 1.8313 | 3.0067 | 0.85654 | 2.4312 | 0.53072 | 1.734
| Rel.Std.Dev. |  | 1.489% | 1.945% | 1.379% | 0.9321% | 3.812% | 1.095% | 1.359% | 1.244% | 1.707% | 2.779% | 1.565% | 1.004% | 0.7614% | 3.353% | 1.802% | 3.593% | 1.228% | 1.52% | 2.375% | 9.161% | 2.257% | 0% | 1.817% | 1.865% | 3.51% | 0.9551% | 2.354% | 0.5254% | 1.753%
| PDM3 |  | -0.08351% | 2.083% | -4.544% | 0.7592% | 4.013% | -2.663% | 17.63% | 0.4324% | 0.4668% | -1.287% | -2.182% | 7.827% | 2.995% | 1.344% | 0.7592% | -9.198% | -1.803% | -0.4275% | -3.179% | -14.7% | -4.603% | -1.975% | 3.356% | 1.309% | -11.61% | -7.461% | 6.555% | 4.233% | 2.083%
| Z-Score (Absolute) |  | 0.13 | 0.43 | 1.28 | 0.08 | 0.92 | 0.8 | 4.44 | 0 | 0.01 | 0.44 | 0.67 | 1.91 | 0.66 | 0.24 | 0.08 | 2.48 | 0.58 | 0.22 | 0.93 | 3.9 | 1.3 | 0.62 | 0.75 | 0.23 | 3.11 | 2.04 | 1.58 | 0.98 | 0.43
| Analytical results for P in OREAS 234b (Certified Value 0.043 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.0431 | 0.0446 | 0.042386 | 0.040906 | 0.04 | 0.0528 | 0.046 | 0.0412 | 0.0431 | 0.042 | 0.045 | 0.045 | 0.047 | 0.042 | 0.0442 | 0.044 | 0.042 | 0.045 | 0.042 | 0.045 | 0.043337 | 0.0435 | 0.0449 | 0.043 | 0.0412 |  |  |  |  | 1
| 1 | 2 | 0.0432 | 0.0464 | 0.042786 | 0.043063 | 0.04 | 0.0528 | 0.046 | 0.044 | 0.043 | 0.043 | 0.044 | 0.043 | 0.044 | 0.041 | 0.0433 | 0.045 | 0.043 | 0.043 | 0.043 | 0.044 | 0.043087 | 0.0435 | 0.0424 | 0.0429 | 0.04 |  |  |  |  | #N/A
| 1 | 3 | 0.0433 | 0.0456 | 0.042534 | 0.044222 | 0.041 | 0.0506 | 0.046 | 0.0427 | 0.0431 | 0.044 | 0.046 | 0.042 | 0.045 | 0.042 | 0.0437 | 0.045 | 0.043 | 0.043 | 0.0445 | 0.045 | 0.043516 | 0.0428 | 0.0426 | 0.0433 | 0.0404 |  |  |  |  | 16
| 1 | 4 | 0.0429 | 0.0441 | 0.042715 | 0.042873 | 0.041 | 0.0511 | 0.045 | 0.0399 | 0.043 | 0.043 | 0.044 | 0.042 | 0.043 | 0.042 | 0.041 | 0.044 | 0.043 | 0.044 | 0.043 | 0.044 | 0.042041 | 0.0444 | 0.043 | 0.0431 | 0.0419 |  |  |  |  | 0.043213
| 1 | 5 | 0.0423 | 0.0455 | 0.043525 | 0.039928 | 0.041 | 0.0537 | 0.046 | 0.0427 | 0.0432 | 0.043 | 0.044 | 0.042 | 0.047 | 0.043 | 0.0438 | 0.046 | 0.041 | 0.045 | 0.043 | 0.046 | 0.042946 | 0.0418 | 0.0431 | 0.043 | 0.043 |  |  |  |  | 49
| 1 | 6 | 0.0438 | 0.0427 | 0.042072 | 0.043096 | 0.038 | 0.0524 | 0.045 | 0.0407 | 0.0428 | 0.042 | 0.044 | 0.043 | 0.045 | 0.044 | 0.0407 | 0.044 | 0.043 | 0.044 | 0.042 | 0.043 | 0.043477 | 0.0438 | 0.0432 | 0.0433 | 0.0424
| Mean |  | 0.0431 | 0.044817 | 0.04267 | 0.042348 | 0.040167 | 0.052233 | 0.045667 | 0.041867 | 0.043033 | 0.042833 | 0.0445 | 0.042833 | 0.045167 | 0.042333 | 0.042783 | 0.044667 | 0.0425 | 0.044 | 0.042917 | 0.0445 | 0.043067 | 0.0433 | 0.0432 | 0.0431 | 0.041483
| Median |  | 0.04315 | 0.04505 | 0.042624 | 0.042968 | 0.0405 | 0.0526 | 0.046 | 0.04195 | 0.04305 | 0.043 | 0.044 | 0.0425 | 0.045 | 0.042 | 0.0435 | 0.0445 | 0.043 | 0.044 | 0.043 | 0.0445 | 0.043212 | 0.0435 | 0.04305 | 0.04305 | 0.04155
| Std Dev. |  | 0.00049396 | 0.0013136 | 0.00049056 | 0.0015997 | 0.001169 | 0.0011639 | 0.0005164 | 0.001524 | 0.00013663 | 0.00075277 | 0.00083666 | 0.001169 | 0.0016021 | 0.0010328 | 0.0015276 | 0.0008165 | 0.00083666 | 0.00089443 | 0.00091742 | 0.0010488 | 0.00054974 | 0.00089889 | 0.00088769 | 0.00016733 | 0.0011635
| Rel.Std.Dev. |  | 1.146% | 2.931% | 1.15% | 3.777% | 2.91% | 2.228% | 1.131% | 3.64% | 0.3175% | 1.757% | 1.88% | 2.729% | 3.547% | 2.44% | 3.571% | 1.828% | 1.969% | 2.033% | 2.138% | 2.357% | 1.276% | 2.076% | 2.055% | 0.3882% | 2.805%
| PDM3 |  | -0.2607% | 3.712% | -1.257% | -2.001% | -7.049% | 20.88% | 5.679% | -3.115% | -0.415% | -0.8778% | 2.979% | -0.8778% | 4.522% | -2.035% | -0.9935% | 3.365% | -1.649% | 1.822% | -0.6849% | 2.979% | -0.3365% | 0.2022% | -0.02926% | -0.2607% | -4.002%
| Z-Score (Absolute) |  | 0.03 | 1.64 | 0.37 | 0.67 | 2.72 | 8.58 | 2.44 | 1.13 | 0.03 | 0.22 | 1.34 | 0.22 | 1.97 | 0.69 | 0.27 | 1.5 | 0.53 | 0.87 | 0.14 | 1.34 | 0 | 0.22 | 0.12 | 0.03 | 1.48
| Analytical results for Pb in OREAS 234b (Certified Value 26.2 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 25.4 | 27.1 | 21.18 | 25.4052 | 25.8394 | 19 | 30.4 | 27 | 25.6 | 27 | 22.3 | 26.7 | 27.6 | 25.5 | 27.4 | 20 | 23 | 24.8 | 27.4 | 33.2 | 27.9932 | 26 | 27.1478 | 26.5 | 99.6871 | 26.7 | 27.3 | 28.8 | 25.1 | 1
| 1 | 2 | 25.4 | 28.3 | 20.85 | 25.7165 | 27.6796 | 25 | 29.4 | 27 | 26.5 | 27 | 24.5 | 25.4 | 26.7 | 25.1 | 26.7 | 20 | 23 | 25.7 | 28 | 24.55 | 26.273 | 26 | 26.5059 | 26.4 | 95.0936 | 26.6 | 26.6 | 28.7 | 25 | 29
| 1 | 3 | 25.1 | 27.6 | 22.4 | 25.8338 | 27.8347 | 20 | 29.6 | 28 | 25.9 | 27 | 23.1 | 26.2 | 27.2 | 24.8 | 27.1 | 19 | 25 | 24.7 | 27.6 | 27.24 | 23.6974 | 26 | 25.4383 | 27.1 | 102.2527 | 26.5 | 25.2 | 27.5 | 26 | 16
| 1 | 4 | 25.3 | 26.9 | 20.74 | 25.6486 | 28.8704 | 18 | 30 | 27 | 26.3 | 27 | 22.8 | 25.3 | 27 | 25 | 25.6 | 18 | 24 | 25.8 | 28.3 | 26.95 | 26.2354 | 27 | 25.6449 | 26.4 | 97.5115 | 26.6 | 27.2 | 27.3 | 27.5 | 26.2337
| 1 | 5 | 24.5 | 27.7 | 21.47 | 24.9899 | 26.3736 | 19 | 30.4 | 28 | 26.1 | 27 | 23.1 | 25.7 | 27.6 | 26.8 | 27.4 | 20 | 25 | 24.4 | 27.2 | 25.47 | 26.1132 | 26 | 26.543 | 24 | 99.5215 | 27.1 | 26.1 | 28 | 26.1 | 50
| 1 | 6 | 25.4 | 26.7 | 21.25 | 23.8038 | 27.8337 | 20 | 30.6 | 27 | 26.1 | 28 | 22.6 | 25.4 | 27.7 | 26.5 | 28.3 | 19 | 24 | 24.5 | 28 | 24.42 | 23.8854 | 27 | 26.6482 | 27.3 | 93.1165 | 26.5 | 26.1 | 27.8 | 28.2
| Mean |  | 25.1833 | 27.3833 | 21.315 | 25.233 | 27.4052 | 20.1667 | 30.0667 | 27.3333 | 26.0833 | 27.1667 | 23.0667 | 25.7833 | 27.3 | 25.6167 | 27.0833 | 19.3333 | 24 | 24.9833 | 27.75 | 26.9717 | 25.6996 | 26.3333 | 26.3213 | 26.2833 | 97.8638 | 26.6667 | 26.4167 | 28.0167 | 26.3167
| Median |  | 25.35 | 27.35 | 21.215 | 25.5269 | 27.7567 | 19.5 | 30.2 | 27 | 26.1 | 27 | 22.95 | 25.55 | 27.4 | 25.3 | 27.25 | 19.5 | 24 | 24.75 | 27.8 | 26.21 | 26.1743 | 26 | 26.5244 | 26.45 | 98.5165 | 26.6 | 26.35 | 27.9 | 26.05
| Std Dev. |  | 0.35449 | 0.5947 | 0.59507 | 0.76142 | 1.1051 | 2.4833 | 0.48442 | 0.5164 | 0.31252 | 0.40825 | 0.76594 | 0.55648 | 0.4 | 0.83766 | 0.89759 | 0.8165 | 0.89443 | 0.61128 | 0.41833 | 3.2709 | 1.6339 | 0.5164 | 0.64968 | 1.1822 | 3.3368 | 0.22509 | 0.78846 | 0.61779 | 1.2891
| Rel.Std.Dev. |  | 1.408% | 2.172% | 2.792% | 3.018% | 4.032% | 12.31% | 1.611% | 1.889% | 1.198% | 1.503% | 3.321% | 2.158% | 1.465% | 3.27% | 3.314% | 4.223% | 3.727% | 2.447% | 1.507% | 12.13% | 6.358% | 1.961% | 2.468% | 4.498% | 3.41% | 0.8441% | 2.985% | 2.205% | 4.898%
| PDM3 |  | -4.004% | 4.382% | -18.75% | -3.815% | 4.466% | -23.13% | 14.61% | 4.192% | -0.573% | 3.557% | -12.07% | -1.717% | 4.065% | -2.352% | 3.239% | -26.3% | -8.514% | -4.766% | 5.78% | 2.813% | -2.036% | 0.38% | 0.3343% | 0.1894% | 273% | 1.651% | 0.6976% | 6.797% | 0.3164%
| Z-Score (Absolute) |  | 0.76 | 0.71 | 3.34 | 0.73 | 0.72 | 4.1 | 2.5 | 0.67 | 0.16 | 0.56 | 2.17 | 0.36 | 0.65 | 0.47 | 0.51 | 4.66 | 1.55 | 0.89 | 0.95 | 0.43 | 0.41 | 0.01 | 0 | 0.03 | 47.67 | 0.23 | 0.06 | 1.13 | 0
| Analytical results for Pr in OREAS 234b (Certified Value 1.75 ppm) | [30 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 1.8 | 1.61 | 1.6196 | 1.7 | 2 | 1.74 | 1.8 | 1.7969 | [21 empty] | 1
| 1 | 2 | 1.9 | 1.66 | 1.6906 | 1.8 | 2 | 1.78 | 1.7 | 1.8406 | [21 empty] | 30
| 1 | 3 | 1.84 | 1.59 | 1.6432 | 1.9 | 2 | 1.76 | 1.9 | 1.7755 | [21 empty] | 16
| 1 | 4 | 1.79 | 1.66 | 1.6546 | 1.7 | 2 | 1.8 | 1.8 | 1.7865 | [21 empty] | 1.7507
| 1 | 5 | 1.82 | 1.69 | 1.6636 | 1.7 | 2 | 1.73 | 1.8 | 1.8666 | [21 empty] | 51
| 1 | 6 | 1.81 | 1.66 | 1.6745 | 1.7 | 2 | 1.8 | 1.8 | 1.8635
| Mean |  | 1.8267 | 1.645 | 1.6577 | 1.75 | 2 | 1.7683 | 1.8 | 1.8216
| Median |  | 1.815 | 1.66 | 1.6591 | 1.7 | 2 | 1.77 | 1.8 | 1.8187
| Std Dev. |  | 0.039833 | 0.037283 | 0.024777 | 0.083666 | 0 | 0.029944 | 0.063246 | 0.04025
| Rel.Std.Dev. |  | 2.181% | 2.266% | 1.495% | 4.781% | 0% | 1.693% | 3.514% | 2.21%
| PDM3 |  | 4.342% | -6.035% | -5.311% | -0.03768% | 14.24% | 1.01% | 2.818% | 4.052%
| Z-Score (Absolute) |  | 0.74 | 1.56 | 1.4 | 0.23 | NA | 0 | 0.4 | 0.67
Results from laboratory 1.08 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Rb in OREAS 234b (Certified Value 15.2 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 14.9 | 15.25 | 13.8887 | 13.4884 | 14.7 | 15.4 | 15.3 | 14.23 | 14.1 | 16.9 | 16.7 | 14 | 15.3 | 14.4 | 15.1 | 14.8 | 15.5 | 15.7197 | 15.3 | 16.2107 | 16.9 | 15.1 | 16.5 | 13.7 | [5 empty] | 1
| 1 | 2 | 15.2 | 14.78 | 13.8498 | 14.0208 | 16 | 15 | 15 | 15.24 | 15.5 | 15.8 | 16.2 | 14 | 15 | 13.8 | 14.7 | 14.6 | 15 | 16.216 | 15.6 | 16.3943 | 16.3 | 14.7 | 16.9 | 13.3 | [5 empty] | 31
| 1 | 3 | 15.8 | 14.53 | 14.0254 | 13.9241 | 15.8 | 14.8 | 15.3 | 14.49 | 13.9 | 16.1 | 16 | 14.2 | 14.9 | 14.2 | 15 | 15.3 | 15 | 16.939 | 15.9 | 16.3038 | 16.4 | 14.7 | 16.1 | 14.1 | [5 empty] | 16
| 1 | 4 | 16.3 | 14.77 | 13.8865 | 13.7848 | 14.7 | 15.4 | 15.4 | 15.06 | 14.7 | 15.4 | 16.5 | 14.1 | 14.7 | 13.8 | 14.9 | 15.5 | 15.5 | 17.6241 | 15.2 | 15.8414 | 16.5 | 15.3 | 16.4 | 15 | [5 empty] | 15.1797
| 1 | 5 | 15.5 | 14.79 | 14.0866 | 13.1614 | 14.7 | 15.4 | 15.7 | 14.7 | 14.3 | 15.3 | 16.6 | 15 | 15.2 | 14.3 | 14.8 | 14.7 | 16 | 15.9114 | 15.6 | 15.8043 | 16.7 | 15.1 | 16.2 | 14.2 | [5 empty] | 52
| 1 | 6 | 16.7 | 14.2 | 14.0362 | 14.2588 | 14.7 | 15.7 | 14.8 | 14.52 | 13.3 | 15.8 | 16.3 | 15.1 | 15.4 | 14.5 | 14.8 | 15.5 | 15.5 | 15.8303 | 14.8 | 15.6956 | 16.4 | 15.4 | 17.4 | 16.1
| Mean |  | 15.7333 | 14.72 | 13.9622 | 13.7731 | 15.1 | 15.2833 | 15.25 | 14.7067 | 14.3 | 15.8833 | 16.3833 | 14.4 | 15.0833 | 14.1667 | 14.8833 | 15.0667 | 15.4167 | 16.3734 | 15.4 | 16.0417 | 16.5333 | 15.05 | 16.5833 | 14.4
| Median |  | 15.65 | 14.775 | 13.9571 | 13.8545 | 14.7 | 15.4 | 15.3 | 14.61 | 14.2 | 15.8 | 16.4 | 14.15 | 15.1 | 14.25 | 14.85 | 15.05 | 15.5 | 16.0637 | 15.45 | 16.026 | 16.45 | 15.1 | 16.45 | 14.15
| Std Dev. |  | 0.67725 | 0.34595 | 0.098688 | 0.39383 | 0.6229 | 0.32506 | 0.31464 | 0.37903 | 0.74833 | 0.57764 | 0.26394 | 0.5099 | 0.26394 | 0.30111 | 0.1472 | 0.41312 | 0.37639 | 0.75426 | 0.38471 | 0.29592 | 0.22509 | 0.29496 | 0.48751 | 1.008
| Rel.Std.Dev. |  | 4.305% | 2.35% | 0.7068% | 2.859% | 4.125% | 2.127% | 2.063% | 2.577% | 5.233% | 3.637% | 1.611% | 3.541% | 1.75% | 2.125% | 0.989% | 2.742% | 2.441% | 4.607% | 2.498% | 1.845% | 1.361% | 1.96% | 2.94% | 7%
| PDM3 |  | 3.647% | -3.029% | -8.021% | -9.267% | -0.5253% | 0.6825% | 0.4629% | -3.116% | -5.795% | 4.635% | 7.929% | -5.137% | -0.6351% | -6.674% | -1.953% | -0.7449% | 1.561% | 7.864% | 1.451% | 5.678% | 8.917% | -0.8547% | 9.247% | -5.137%
| Z-Score (Absolute) |  | 0.65 | 0.38 | 1.14 | 1.33 | 0.01 | 0.19 | 0.16 | 0.39 | 0.8 | 0.8 | 1.31 | 0.7 | 0.01 | 0.94 | 0.21 | 0.03 | 0.33 | 1.3 | 0.31 | 0.96 | 1.46 | 0.04 | 1.51 | 0.7
| Analytical results for Re in OREAS 234b (Indicative Value 0.003 ppm) | [30 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.29 | 1.30 | 1.31 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 | 0.003 | < 0.01 | 0.003 | < 0.05 | 0.003 | < 0.005 | 0.003 | 0.003 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0184 | 0.005 | 0.003 | 0.004 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.01 | 0.003 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.003 | 0.002 | < 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | 0.0173 | 0.004 | 0.003 | 0.004 | [10 empty] | 3
| 1 | 3 | < 0.002 | < 0.01 | 0.003 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0172 | 0.004 | 0.003 | 0.004 | [10 empty] | 16
| 1 | 4 | < 0.002 | < 0.01 | 0.003 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.002 | 0.003 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0145 | 0.005 | 0.003 | 0.006 | [10 empty] | 0.0030606
| 1 | 5 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.005 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0164 | 0.006 | 0.003 | 0.006 | [10 empty] | 9
| 1 | 6 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.002 | 0.002 | 0.002 | < 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0173 | 0.005 | 0.005 | 0.006
| Mean |  | 0.003 |  | 0.0026667 |  | 0.0036667 |  | 0.0026667 | 0.0023333 | 0.0021667 | 0.002 |  |  |  |  | 0.002 | 0.01685 | 0.0048333 | 0.0033333 | 0.005
| Median |  | 0.003 |  | 0.003 |  | 0.003 |  | 0.003 | 0.002 | 0.002 | 0.002 |  |  |  |  | 0.002 | 0.01725 | 0.005 | 0.003 | 0.005
| Std Dev. |  |  |  | 0.0005164 |  | 0.0011547 |  | 0.0005164 | 0.0005164 | 0.00040825 | 0 |  |  |  |  | 0 | 0.0013157 | 0.00075277 | 0.0008165 | 0.0010954
| Rel.Std.Dev. |  |  |  | 19.36% |  | 31.49% |  | 19.36% | 22.13% | 18.84% | 0% |  |  |  |  | 0% | 7.808% | 15.57% | 24.49% | 21.91%
| PDM3 |  | -1.98% |  | -12.87% |  | 19.8% |  | -12.87% | -23.76% | -29.21% | -34.65% |  |  |  |  | -34.65% | 450.5% | 57.92% | 8.911% | 63.37%
| Z-Score (Absolute) |  | 0.9 | 1.57 | 0 | 15.06 | 0.22 | 0.11 | 0 | 0.22 | 0.34 | 0.67 | 1683.97 | 15.06 | 0.11 | 15.06 | 0.56 | 9.56 | 1.46 | 0.45 | 1.57
| Analytical results for S in OREAS 234b (Certified Value 0.405 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.42 | 0.4 | 0.41718 | 0.40386 | 0.39 | 0.463 | 0.41 | 0.41 | 0.41 | 0.37 | 0.42 | 0.42 | 0.41 | 0.4 | 0.4 | 0.406 | 0.4 | 0.431 | 0.38028 | 0.4 | 0.39 | 0.37221 | 0.43 | 0.42 | 0.41 | 0.42 |  |  |  | 1
| 1 | 2 | 0.42 | 0.43 | 0.41745 | 0.40828 | 0.38 | 0.451 | 0.401 | 0.38 | 0.39 | 0.37 | 0.4 | 0.41 | 0.39 | 0.39 | 0.4 | 0.416 | 0.4 | 0.415 | 0.38475 | 0.405 | 0.39 | 0.37484 | 0.42 | 0.4 | 0.41 | 0.4 |  |  |  | 17
| 1 | 3 | 0.41 | 0.41 | 0.41798 | 0.41849 | 0.38 | 0.453 | 0.405 | 0.4 | 0.4 | 0.39 | 0.41 | 0.4 | 0.4 | 0.4 | 0.4 | 0.409 | 0.4 | 0.416 | 0.36933 | 0.405 | 0.39 | 0.37107 | 0.43 | 0.4 | 0.41 | 0.41 |  |  |  | 16
| 1 | 4 | 0.41 | 0.4 | 0.41769 | 0.40196 | 0.38 | 0.455 | 0.409 | 0.38 | 0.4 | 0.39 | 0.4 | 0.4 | 0.39 | 0.39 | 0.38 | 0.413 | 0.4 | 0.419 | 0.38458 | 0.405 | 0.4 | 0.36386 | 0.42 | 0.4 | 0.41 | 0.4 |  |  |  | 0.40505
| 1 | 5 | 0.4 | 0.41 | 0.41254 | 0.39823 | 0.38 | 0.465 | 0.41 | 0.37 | 0.41 | 0.37 | 0.4 | 0.4 | 0.41 | 0.4 | 0.4 | 0.417 | 0.4 | 0.431 | 0.37641 | 0.4 | 0.4 | 0.37978 | 0.44 | 0.41 | 0.41 | 0.41 |  |  |  | 53
| 1 | 6 | 0.41 | 0.4 | 0.41816 | 0.41658 | 0.38 | 0.452 | 0.4 | 0.4 | 0.4 | 0.38 | 0.4 | 0.41 | 0.4 | 0.4 | 0.37 | 0.403 | 0.4 | 0.422 | 0.37156 | 0.41 | 0.38 | 0.37157 | 0.43 | 0.4 | 0.41 | 0.4
| Mean |  | 0.41167 | 0.40833 | 0.41683 | 0.4079 | 0.38167 | 0.4565 | 0.40583 | 0.39 | 0.40167 | 0.37833 | 0.405 | 0.40667 | 0.4 | 0.39667 | 0.39167 | 0.41067 | 0.4 | 0.42233 | 0.37782 | 0.40417 | 0.39167 | 0.37222 | 0.42833 | 0.405 | 0.41 | 0.40667
| Median |  | 0.41 | 0.405 | 0.41757 | 0.40607 | 0.38 | 0.454 | 0.407 | 0.39 | 0.4 | 0.375 | 0.4 | 0.405 | 0.4 | 0.4 | 0.4 | 0.411 | 0.4 | 0.4205 | 0.37834 | 0.405 | 0.39 | 0.37189 | 0.43 | 0.4 | 0.41 | 0.405
| Std Dev. |  | 0.0075277 | 0.01169 | 0.002133 | 0.0081574 | 0.0040825 | 0.0059917 | 0.004535 | 0.015492 | 0.0075277 | 0.0098319 | 0.0083666 | 0.008165 | 0.0089443 | 0.005164 | 0.013292 | 0.0056095 | 6.0809e-17 | 0.0071461 | 0.0065269 | 0.0037639 | 0.0075277 | 0.005208 | 0.0075277 | 0.0083666 | 6.0809e-17 | 0.008165
| Rel.Std.Dev. |  | 1.829% | 2.863% | 0.5117% | 2% | 1.07% | 1.313% | 1.117% | 3.972% | 1.874% | 2.599% | 2.066% | 2.008% | 2.236% | 1.302% | 3.394% | 1.366% | 1.52e-14% | 1.692% | 1.728% | 0.9313% | 1.922% | 1.399% | 1.757% | 2.066% | 1.483e-14% | 2.008%
| PDM3 |  | 1.633% | 0.8098% | 2.908% | 0.7029% | -5.774% | 12.7% | 0.1926% | -3.716% | -0.8361% | -6.597% | -0.01316% | 0.3983% | -1.248% | -2.071% | -3.305% | 1.386% | -1.248% | 4.266% | -6.724% | -0.2189% | -3.305% | -8.106% | 5.747% | -0.01316% | 1.221% | 0.3983%
| Z-Score (Absolute) |  | 0.73 | 0.36 | 1.29 | 0.32 | 2.55 | 5.63 | 0.09 | 1.64 | 0.36 | 2.92 | 0 | 0.18 | 0.55 | 0.91 | 1.46 | 0.62 | 0.55 | 1.9 | 2.97 | 0.09 | 1.46 | 3.58 | 2.55 | 0 | 0.55 | 0.18
| Analytical results for Sb in OREAS 234b (Certified Value 1.43 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.62 | 1.47 | 1.4283 | 0.9 | 1.4 | 1.42 | < 5 | 1.7 | 1.47 | 1.45 | 1.36 | 1.5 | < 5 | 1.2 | 1.3 | 1.37 | 6.13 | 1.5 | 1.3837 | 1.43 | 1.5014 | 1.47 | 1.51 | 1.6 | 1.34 |  |  |  |  | 1
| 1 | 2 | 1.62 | 1.5 | 1.4909 | 1.3 | 1.4 | 1.5 | < 5 | 1.7 | 1.37 | 1.47 | 1.4 | 1.49 | < 5 | 1.3 | 1.3 | 1.39 | 5.76 | 1.5 | 1.3698 | 1.4 | 1.4844 | 1.45 | 1.46 | 1.6 | 1.29 |  |  |  |  | 33
| 1 | 3 | 1.59 | 1.55 | 1.6151 | 1.5 | 1.4 | 1.37 | < 5 | 1.6 | 1.42 | 1.51 | 1.31 | 1.52 | < 5 | 1.2 | 1.3 | 1.39 | 7.71 | 1.5 | 1.406 | 1.3 | 1.4346 | 1.45 | 1.49 | 1.5 | 1.39 |  |  |  |  | 16
| 1 | 4 | 1.65 | 1.44 | 1.4773 | 1.3 | 1.5 | 1.4 | < 5 | 1.5 | 1.39 | 1.44 | 1.37 | 1.44 | < 5 | 1.2 | 1.3 | 1.4 | 7.83 | 1.5 | 1.4179 | 1.48 | 1.4166 | 1.38 | 1.59 | 1.5 | 1.48 |  |  |  |  | 1.4269
| 1 | 5 | 1.76 | 1.51 | 1.3902 | 1.3 | 1.4 | 1.47 | < 5 | 1.6 | 1.38 | 1.48 | 1.46 | 1.5 | < 5 | 1.3 | 1.3 | 1.3 | 8.61 | 1.5 | 1.3733 | 1.43 | 1.4447 | 1.44 | 1.64 | 1.5 | 1.4 |  |  |  |  | 54
| 1 | 6 | 1.67 | 1.45 | 1.4881 | 0.1 | 1.5 | 1.32 | < 5 | 1.7 | 1.4 | 1.5 | 1.49 | 1.47 | < 5 | 1.3 | 1.3 | 1.36 | 7.8 | 1.5 | 1.4134 | 1.48 | 1.4258 | 1.38 | 1.5 | 1.9 | 1.57
| Mean |  | 1.6517 | 1.4867 | 1.4817 | 1.0667 | 1.4333 | 1.4133 |  | 1.6333 | 1.405 | 1.475 | 1.3983 | 1.4867 |  | 1.25 | 1.3 | 1.3683 | 7.3067 | 1.5 | 1.394 | 1.42 | 1.4513 | 1.4283 | 1.5317 | 1.6 | 1.4117
| Median |  | 1.635 | 1.485 | 1.4827 | 1.3 | 1.4 | 1.41 |  | 1.65 | 1.395 | 1.475 | 1.385 | 1.495 |  | 1.25 | 1.3 | 1.38 | 7.755 | 1.5 | 1.3948 | 1.43 | 1.4397 | 1.445 | 1.505 | 1.55 | 1.395
| Std Dev. |  | 0.059805 | 0.041312 | 0.07635 | 0.51251 | 0.05164 | 0.065625 |  | 0.08165 | 0.036194 | 0.027386 | 0.066758 | 0.028048 |  | 0.054772 | 0 | 0.03656 | 1.1095 | 0 | 0.021014 | 0.066633 | 0.034007 | 0.038687 | 0.068532 | 0.15492 | 0.10028
| Rel.Std.Dev. |  | 3.621% | 2.779% | 5.153% | 48.05% | 3.603% | 4.643% |  | 4.999% | 2.576% | 1.857% | 4.774% | 1.887% |  | 4.382% | 0% | 2.672% | 15.19% | 0% | 1.507% | 4.692% | 2.343% | 2.709% | 4.474% | 9.682% | 7.104%
| PDM3 |  | 15.75% | 4.185% | 3.835% | -25.25% | 0.4476% | -0.954% |  | 14.46% | -1.538% | 3.368% | -2.005% | 4.185% |  | -12.4% | -8.896% | -4.108% | 412% | 5.12% | -2.307% | -0.4868% | 1.703% | 0.09715% | 7.339% | 12.13% | -1.071%
| Z-Score (Absolute) |  | 2.55 | 0.45 | 0.39 | 4.9 | 0.23 | 0.48 | 13.36 | 2.32 | 0.59 | 0.3 | 0.67 | 0.45 | 13.36 | 2.56 | 1.93 | 1.06 | 74.61 | 0.62 | 0.73 | 0.4 | 0 | 0.29 | 1.02 | 1.9 | 0.5
| Analytical results for Sc in OREAS 234b (Certified Value 36.8 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 39.5 | 36.6 | 32.77 | 36.528 | 34.8755 | 36 | 42.4 | 36 | 35.3 | 36 | 32.4 | 39.9 | 38.5 | 36.9 | 37.7 | 36 | 31 | 36 | 35.4 | 42 | 38.1567 | 37.7 | 34.9762 | 38.7 | 38.7 | 39.6 | 34.7 |  |  | 1
| 1 | 2 | 39.7 | 37.4 | 33.47 | 36.5952 | 36.6977 | 37 | 39.6 | 36 | 38.1 | 35 | 34.3 | 38.4 | 38.1 | 35.9 | 36.1 | 36 | 33 | 35 | 34.6 | 39 | 38.2779 | 38.6 | 34.5472 | 37.9 | 37 | 39.2 | 34.8 |  |  | 34
| 1 | 3 | 39.5 | 37.3 | 30.47 | 36.7008 | 37.6392 | 36 | 40.1 | 36 | 36.5 | 35 | 33.5 | 39.2 | 39 | 36.4 | 35 | 36 | 31 | 35 | 37.6 | 39 | 38.6568 | 38.4 | 34.9663 | 37.2 | 37 | 38 | 36.2 |  |  | 16
| 1 | 4 | 40 | 37 | 32.1 | 36.432 | 36.3757 | 36 | 40.9 | 35 | 38.3 | 35 | 34.6 | 38.1 | 38.9 | 35.9 | 35.5 | 34 | 33 | 36 | 36.5 | 40 | 39.1235 | 37.9 | 34.0892 | 38.7 | 37.8 | 37.2 | 38.2 |  |  | 36.7547
| 1 | 5 | 39.6 | 36.3 | 32.68 | 36.768 | 34.0812 | 37 | 41.2 | 36 | 37.2 | 36 | 34.3 | 37.7 | 38.2 | 38.2 | 35.4 | 36 | 34 | 36 | 36.6 | 41 | 38.8146 | 35.2 | 35.1762 | 38 | 38 | 38.1 | 36 |  |  | 55
| 1 | 6 | 40.6 | 35.8 | 32.05 | 37.3248 | 36.4442 | 37 | 40.9 | 35 | 37.6 | 34 | 33.4 | 38.3 | 38.9 | 40.3 | 37.3 | 33 | 34 | 35 | 36.9 | 40 | 38.502 | 36.4 | 34.9166 | 37.6 | 37.9 | 37.5 | 40.6
| Mean |  | 39.8167 | 36.7333 | 32.2567 | 36.7248 | 36.0189 | 36.5 | 40.85 | 35.6667 | 37.1667 | 35.1667 | 33.75 | 38.6 | 38.6 | 37.2667 | 36.1667 | 35.1667 | 32.6667 | 35.5 | 36.2667 | 40.1667 | 38.5886 | 37.3667 | 34.7786 | 38.0167 | 37.7333 | 38.2667 | 36.75
| Median |  | 39.65 | 36.8 | 32.39 | 36.648 | 36.41 | 36.5 | 40.9 | 36 | 37.4 | 35 | 33.9 | 38.35 | 38.7 | 36.65 | 35.8 | 36 | 33 | 35.5 | 36.55 | 40 | 38.5794 | 37.8 | 34.9415 | 37.95 | 37.85 | 38.05 | 36.1
| Std Dev. |  | 0.42622 | 0.6186 | 1.0173 | 0.3174 | 1.3005 | 0.54772 | 0.96488 | 0.5164 | 1.1201 | 0.75277 | 0.8167 | 0.80498 | 0.38987 | 1.7143 | 1.0985 | 1.3292 | 1.3663 | 0.54772 | 1.0838 | 1.169 | 0.35564 | 1.3125 | 0.39503 | 0.59805 | 0.65013 | 0.94587 | 2.2713
| Rel.Std.Dev. |  | 1.07% | 1.684% | 3.154% | 0.8643% | 3.611% | 1.501% | 2.362% | 1.448% | 3.014% | 2.141% | 2.42% | 2.085% | 1.01% | 4.6% | 3.037% | 3.78% | 4.182% | 1.543% | 2.988% | 2.91% | 0.9216% | 3.512% | 1.136% | 1.573% | 1.723% | 2.472% | 6.181%
| PDM3 |  | 8.331% | -0.0582% | -12.24% | -0.08142% | -2.002% | -0.693% | 11.14% | -2.96% | 1.121% | -4.321% | -8.175% | 5.021% | 5.021% | 1.393% | -1.6% | -4.321% | -11.12% | -3.414% | -1.328% | 9.283% | 4.989% | 1.665% | -5.376% | 3.433% | 2.663% | 4.114% | -0.01286%
| Z-Score (Absolute) |  | 1.62 | 0 | 2.35 | 0 | 0.38 | 0.12 | 2.16 | 0.56 | 0.23 | 0.82 | 1.57 | 0.98 | 0.98 | 0.28 | 0.3 | 0.82 | 2.14 | 0.65 | 0.25 | 1.8 | 0.97 | 0.33 | 1.03 | 0.67 | 0.53 | 0.81 | 0.01
| Analytical results for Se in OREAS 234b (Indicative Value 0.99 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 2 | 0.5 | < 1.5 | 0.4 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 1.59 | < 5 | 1 | 0.586 | 3.2 | < 2 | < 1 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 2 | 0.5 | < 1.5 | 0.7 | < 1 | 0.6 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 4.96 | < 5 | 1 | 0.611 | 1.9 | < 2 | < 1 | < 0.5 | [6 empty] | 4
| 1 | 3 | < 2 | < 0.5 | < 1.5 | 0.6 | < 1 | < 0.5 | < 5 | 2 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 6.13 | < 5 | 1 | 0.5816 | 2.5 | < 2 | < 1 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 2 | 0.6 | < 1.5 | 0.5 | < 1 | < 0.5 | < 5 | 2 | 1 | < 1 | 1 | < 5 | < 2 | < 1 | 1 | 6.21 | < 5 | 1 | 0.5558 | 2.3 | < 2 | < 1 | < 0.5 | [6 empty] | 0.98633
| 1 | 5 | < 2 | 0.6 | < 1.5 | 0.5 | < 1 | 0.5 | < 5 | 1 | 1 | < 1 | 1 | < 5 | < 2 | < 1 | 1 | 4.07 | < 5 | 1 | 0.5643 | 1.7 | < 2 | < 1 | < 0.5 | [6 empty] | 10
| 1 | 6 | < 2 | 0.6 | < 1.5 | 0.4 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.54 | < 5 | 1 | 0.5574 | 2.6 | < 2 | < 1 | 0.6
| Mean |  |  | 0.56 |  | 0.51667 |  | 0.55 |  | 1.3333 | 1 | 1 | 1.5 |  |  |  | 1 | 4.9167 |  | 1 | 0.57602 | 2.3667 |  |  | 0.6
| Median |  |  | 0.6 |  | 0.5 |  | 0.55 |  | 1 | 1 | 1 | 1.5 |  |  |  | 1 | 5.545 |  | 1 | 0.57295 | 2.4 |  |  | 0.6
| Std Dev. |  |  | 0.054772 |  | 0.1169 |  | 0.070711 |  | 0.5164 | 0 | 0 | 0.54772 |  |  |  | 0 | 1.8743 |  | 0 | 0.0212 | 0.53541
| Rel.Std.Dev. |  |  | 9.781% |  | 22.63% |  | 12.86% |  | 38.73% | 0% | 0% | 36.51% |  |  |  | 0% | 38.12% |  | 0% | 3.68% | 22.62%
| PDM3 |  |  | -43.22% |  | -47.62% |  | -44.24% |  | 35.18% | 1.385% | 1.385% | 52.08% |  |  |  | 1.385% | 398.5% |  | 1.385% | -41.6% | 139.9% |  |  | -39.17%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.34 | 0.66 | 0.69 | 0.89 | 2.06 | 0.46 | 0 | 0.23 | 0.69 | 2.06 | 0 | 0.69 | 0 | 5.37 | 2.06 | 0 | 0.58 | 1.87 | 0 | 0.69 | 0.95
| Analytical results for Sm in OREAS 234b (Certified Value 2.48 ppm) | [30 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.43 | 2 | 2.2346 | 2.4 | 2.6 | 2.59 | 2.4 | 2.472 | [21 empty] | 1
| 1 | 2 | 2.59 | 2.1 | 2.2241 | 3.1 | 2.4 | 2.59 | 2.5 | 2.5282 | [21 empty] | 19
| 1 | 3 | 2.52 | 1.99 | 2.2379 | 2.8 | 2.4 | 2.59 | 2.4 | 2.439 | [21 empty] | 16
| 1 | 4 | 2.41 | 2.03 | 2.3061 | 2.3 | 2.5 | 2.63 | 2.4 | 2.4101 | [21 empty] | 2.4849
| 1 | 5 | 2.55 | 2.13 | 2.2576 | 2.3 | 2.5 | 2.52 | 2.5 | 2.5592 | [21 empty] | 56
| 1 | 6 | 2.38 | 2.03 | 2.2324 | 2.3 | 2.4 | 2.66 | 2.5 | 2.5686
| Mean |  | 2.48 | 2.0467 | 2.2488 | 2.5333 | 2.4667 | 2.5967 | 2.45 | 2.4962
| Median |  | 2.475 | 2.03 | 2.2363 | 2.35 | 2.45 | 2.59 | 2.45 | 2.5001
| Std Dev. |  | 0.084853 | 0.056095 | 0.030187 | 0.33862 | 0.08165 | 0.047188 | 0.054772 | 0.065568
| Rel.Std.Dev. |  | 3.421% | 2.741% | 1.342% | 13.37% | 3.31% | 1.817% | 2.236% | 2.627%
| PDM3 |  | -0.198% | -17.64% | -9.503% | 1.948% | -0.7346% | 4.497% | -1.405% | 0.4535%
| Z-Score (Absolute) |  | 0.11 | 6.9 | 3.63 | 0.97 | 0.11 | 2 | 0.38 | 0.37
| Analytical results for Sn in OREAS 234b (Certified Value 1.14 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.5 | 1.2 | 4.34 | < 2.5 | 1 | 1.33 | 1.5 | 1.1 | < 10 | 1.3 | 1.1 | 1.2 | 1 | 1.1 | < 10 | 0.9 | 1 | 1.1 | 2.25 | 1.5 | 1.0207 | 1 | 0.8025 | 1.4 | 1.3 | 1.3 | 1.6 |  |  | 1
| 1 | 2 | 1.5 | 1.2 | 1.98 | < 2.5 | 1 | 1.23 | 1.5 | 1.1 | < 10 | 1.3 | 1 | 1.2 | 1 | 1.1 | < 10 | 1 | 1.1 | 1.1 | 2.01 | 1 | 1.0502 | 1.1 | 0.7644 | 1.4 | 1.2 | 1.3 | 1.6 |  |  | 20
| 1 | 3 | 1.4 | 1.2 | 1.89 | < 2.5 | 1 | 1.22 | 1.6 | 1.1 | < 10 | 1.3 | 1 | 1.1 | 1.1 | 1.1 | < 10 | 0.9 | 1 | 1.1 | 2.32 | 1.5 | 1.0551 | 1.1 | 0.7746 | 1.3 | 1.2 | 1.3 | 1.7 |  |  | 16
| 1 | 4 | 1.4 | 1.1 | 4.6 | < 2.5 | 1 | 1.26 | 1.7 | 1.1 | < 10 | 1.2 | 1 | 1.1 | 1 | 1.1 | < 10 | 0.9 | 1.1 | 1.1 | 2.13 | 1.5 | 1.0507 | 1.1 | 0.7779 | 1.2 | 1.1 | 1.3 | 2 |  |  | 1.1388
| 1 | 5 | 1.4 | 1.1 | 6.42 | < 2.5 | 1 | 1.27 | 1.6 | 1.1 | < 10 | 1.3 | 1 | 1.1 | 1.1 | 1.1 | < 10 | 1 | 0.9 | 1 | 2.18 | 1 | 1.066 | 1.1 | 0.7351 | 1.3 | 1.2 | 1.3 | 1.9 |  |  | 57
| 1 | 6 | 1.5 | 1.1 | 2.53 | < 2.5 | < 1 | 1.24 | 1.8 | 1.1 | < 10 | 1.3 | 1 | 1.1 | 1.2 | 1.1 | < 10 | 0.9 | 1.2 | 1.1 | 1.47 | 1 | 1.08 | 1.2 | 0.7952 | 1.3 | 1.2 | 1.4 | 2.1
| Mean |  | 1.45 | 1.15 | 3.6267 |  | 1 | 1.2583 | 1.6167 | 1.1 |  | 1.2833 | 1.0167 | 1.1333 | 1.0667 | 1.1 |  | 0.93333 | 1.05 | 1.0833 | 2.06 | 1.25 | 1.0538 | 1.1 | 0.77495 | 1.3167 | 1.2 | 1.3167 | 1.8167
| Median |  | 1.45 | 1.15 | 3.435 |  | 1 | 1.25 | 1.6 | 1.1 |  | 1.3 | 1 | 1.1 | 1.05 | 1.1 |  | 0.9 | 1.05 | 1.1 | 2.155 | 1.25 | 1.0529 | 1.1 | 0.77625 | 1.3 | 1.2 | 1.3 | 1.8
| Std Dev. |  | 0.054772 | 0.054772 | 1.7994 |  | 0 | 0.039707 | 0.1169 | 2.4324e-16 |  | 0.040825 | 0.040825 | 0.05164 | 0.08165 | 2.4324e-16 |  | 0.05164 | 0.10488 | 0.040825 | 0.30777 | 0.27386 | 0.019798 | 0.063246 | 0.023988 | 0.075277 | 0.063246 | 0.040825 | 0.2137
| Rel.Std.Dev. |  | 3.777% | 4.763% | 49.62% |  | 0% | 3.156% | 7.231% | 2.211e-14% |  | 3.181% | 4.016% | 4.556% | 7.655% | 2.211e-14% |  | 5.533% | 9.989% | 3.768% | 14.94% | 21.91% | 1.879% | 5.75% | 3.095% | 5.717% | 5.27% | 3.101% | 11.76%
| PDM3 |  | 27.33% | 0.9841% | 218.5% |  | -12.19% | 10.5% | 41.96% | -3.407% |  | 12.69% | -10.72% | -0.4794% | -6.334% | -3.407% |  | -18.04% | -7.797% | -4.87% | 80.89% | 9.765% | -7.465% | -3.407% | -31.95% | 15.62% | 5.375% | 15.62% | 59.53%
| Z-Score (Absolute) |  | 1.01 | 0.34 | 10.8 | 0.11 | NA | 0.15 | 1.76 | 0.56 | 16.97 | 0.26 | 0.94 | 0.41 | 0.71 | 0.56 | 16.97 | 1.31 | 0.79 | 0.64 | 3.75 | 0.11 | 0.77 | 0.56 | 2.02 | 0.41 | 0.11 | 0.41 | 2.66
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 234b (Certified Value 194 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 183 | 192.94 | 202 | 193.3932 | 185.0243 | 181 | 197 | 199 | 185.57 | 186 | 180 | 208 | 207 | 206 | 203 | 210 | 189.6 | 187 | 196 | 140.04 | 187.435 | 184 | 187.8873 | 197 | 214.685 | 196 | 206.1 | 196.1 | 175 | 1
| 1 | 2 | 178.5 | 189.01 | 198 | 195.0396 | 197.8662 | 201 | 201 | 195 | 193.47 | 182 | 182 | 202 | 204 | 196.5 | 200 | 207 | 187.1 | 180 | 189 | 116.51 | 180.899 | 188 | 196.5208 | 199 | 217.8711 | 191 | 191.7 | 195.3 | 179 | 21
| 1 | 3 | 181.3 | 190.83 | 200 | 195.1278 | 203.7429 | 199 | 198 | 203 | 190.79 | 183 | 186 | 208 | 193.5 | 196.5 | 202 | 206 | 187.4 | 180 | 187 | 120.2 | 177.4695 | 187 | 197.5875 | 206 | 221.8835 | 194 | 194.2 | 197.5 | 189 | 16
| 1 | 4 | 182.7 | 194.14 | 196 | 193.6774 | 194.6748 | 179 | 193 | 194 | 192.25 | 183 | 185 | 202 | 196 | 192.5 | 199.5 | 196 | 189.7 | 182 | 189 | 120.68 | 181.2885 | 184 | 198.1657 | 196.5 | 212.4084 | 192 | 205 | 196.6 | 204 | 193.6714
| 1 | 5 | 173.2 | 186.46 | 185 | 193.9616 | 183.5092 | 183 | 198 | 200 | 191.18 | 183 | 184 | 199.5 | 201 | 210 | 204 | 204 | 187.1 | 177 | 199 | 127.01 | 182.2385 | 187 | 195.4164 | 201 | 210.913 | 195 | 205.1 | 195.4 | 192 | 58
| 1 | 6 | 185.1 | 184.26 | 187 | 197.6268 | 195.1149 | 184 | 199 | 200 | 190.93 | 184 | 183 | 205 | 200 | 198 | 205 | 191 | 188.9 | 178 | 193 | 105.42 | 176.9565 | 187 | 194.3286 | 195.5 | 204.4499 | 194 | 202.7 | 196.8 | 217
| Mean |  | 180.6333 | 189.6067 | 194.6667 | 194.8044 | 193.322 | 187.8333 | 197.6667 | 198.5 | 190.6983 | 183.5 | 183.3333 | 204.0833 | 200.25 | 199.9167 | 202.25 | 202.3333 | 188.3 | 180.6667 | 192.1667 | 121.6433 | 181.0478 | 186.1667 | 194.9844 | 199.1667 | 213.7018 | 193.6667 | 200.8 | 196.2833 | 192.6667
| Median |  | 182 | 189.92 | 197 | 194.5006 | 194.8949 | 183.5 | 198 | 199.5 | 191.055 | 183 | 183.5 | 203.5 | 200.5 | 197.25 | 202.5 | 205 | 188.15 | 180 | 191 | 120.44 | 181.0938 | 187 | 195.9686 | 198 | 213.5467 | 194 | 203.85 | 196.35 | 190.5
| Std Dev. |  | 4.2444 | 3.7948 | 7.0333 | 1.5555 | 7.7386 | 9.6003 | 2.6583 | 3.3912 | 2.7086 | 1.3784 | 2.1602 | 3.4988 | 4.9774 | 6.6439 | 2.1852 | 7.2847 | 1.241 | 3.559 | 4.6655 | 11.4874 | 3.7904 | 1.7224 | 3.7481 | 3.8816 | 6.0056 | 1.8619 | 6.2322 | 0.85186 | 15.7056
| Rel.Std.Dev. |  | 2.35% | 2.001% | 3.613% | 0.7985% | 4.003% | 5.111% | 1.345% | 1.708% | 1.42% | 0.7512% | 1.178% | 1.714% | 2.486% | 3.323% | 1.08% | 3.6% | 0.659% | 1.97% | 2.428% | 9.443% | 2.094% | 0.9252% | 1.922% | 1.949% | 2.81% | 0.9614% | 3.104% | 0.434% | 8.152%
| PDM3 |  | -6.732% | -2.099% | 0.5139% | 0.585% | -0.1804% | -3.014% | 2.063% | 2.493% | -1.535% | -5.252% | -5.338% | 5.376% | 3.397% | 3.225% | 4.429% | 4.473% | -2.773% | -6.715% | -0.7769% | -37.19% | -6.518% | -3.875% | 0.678% | 2.837% | 10.34% | -0.002432% | 3.681% | 1.349% | -0.5188%
| Z-Score (Absolute) |  | 1.51 | 0.47 | 0.12 | 0.13 | 0.04 | 0.67 | 0.46 | 0.56 | 0.34 | 1.18 | 1.19 | 1.2 | 0.76 | 0.72 | 0.99 | 1 | 0.62 | 1.5 | 0.17 | 8.33 | 1.46 | 0.87 | 0.15 | 0.64 | 2.32 | 0 | 0.82 | 0.3 | 0.12
| Analytical results for Ta in OREAS 234b (Certified Value 0.24 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.27 | 0.24 | 0.39917 | < 0.1 | 0.23 | 0.37 | 0.22 | 12 | 0.4 | 0.23 | 0.25 | 0.2 | 0.29 | 0.18 | 0.2 | 0.22 | < 0.5 | 0.19196 | 0.24 | 0.3505 | 0.99 | 0.75 | 0.32 | 0.21 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.24 | 0.41629 | 0.2 | 0.2 | 0.35 | 0.23 | 12 | 0.4 | 0.21 | 0.24 | 0.22 | 0.26 | 0.2 | 0.2 | 0.23 | < 0.5 | 0.19226 | 0.23 | 0.3262 | 0.49 | 0.77 | 0.31 | 0.21 | [5 empty] | 22
| 1 | 3 | 0.23 | 0.24 | 0.50124 | 0.2 | 0.2 | 0.4 | 0.22 | 12 | 0.4 | 0.22 | 0.23 | 0.21 | 0.27 | 0.2 | 0.2 | 0.23 | < 0.5 | 0.18752 | 0.24 | 0.3332 | 0.52 | 0.7 | 0.3 | 0.22 | [5 empty] | 16
| 1 | 4 | 0.23 | 0.24 | 0.38481 | 0.2 | 0.21 | 0.38 | 0.23 | 12 | 0.4 | 0.22 | 0.24 | 0.21 | 0.3 | 0.18 | 0.2 | 0.23 | < 0.5 | 0.18362 | 0.24 | 0.3228 | 0.58 | 0.69 | 0.33 | 0.24 | [5 empty] | 0.2386
| 1 | 5 | 0.22 | 0.24 | 0.31976 | 0.2 | 0.2 | 0.38 | 0.23 | 12 | 0.4 | 0.22 | 0.25 | 0.23 | 0.27 | 0.19 | 0.2 | 0.22 | < 0.5 | 0.19438 | 0.21 | 0.3481 | 0.46 | 0.74 | 0.32 | 0.23 | [5 empty] | 59
| 1 | 6 | 0.17 | 0.24 | 0.42817 | < 0.1 | 0.2 | 0.39 | 0.22 | 12 | 0.3 | 0.22 | 0.25 | 0.23 | 0.29 | 0.19 | 0.2 | 0.22 | < 0.5 | 0.19508 | 0.22 | 0.4104 | 0.44 | 0.79 | 0.32 | 0.25
| Mean |  | 0.23667 | 0.24 | 0.40824 | 0.2 | 0.20667 | 0.37833 | 0.225 | 12 | 0.38333 | 0.22 | 0.24333 | 0.21667 | 0.28 | 0.19 | 0.2 | 0.225 |  | 0.1908 | 0.23 | 0.34853 | 0.58 | 0.74 | 0.31667 | 0.22667
| Median |  | 0.23 | 0.24 | 0.40773 | 0.2 | 0.2 | 0.38 | 0.225 | 12 | 0.4 | 0.22 | 0.245 | 0.215 | 0.28 | 0.19 | 0.2 | 0.225 |  | 0.19211 | 0.235 | 0.34065 | 0.505 | 0.745 | 0.32 | 0.225
| Std Dev. |  | 0.044572 | 0 | 0.059282 | 0 | 0.012111 | 0.017224 | 0.0054772 | 0 | 0.040825 | 0.0063246 | 0.008165 | 0.012111 | 0.015492 | 0.0089443 | 3.0405e-17 | 0.0054772 |  | 0.0044028 | 0.012649 | 0.032335 | 0.20678 | 0.038987 | 0.010328 | 0.01633
| Rel.Std.Dev. |  | 18.83% | 0% | 14.52% | 0% | 5.86% | 4.553% | 2.434% | 0% | 10.65% | 2.875% | 3.355% | 5.59% | 5.533% | 4.708% | 1.52e-14% | 2.434% |  | 2.308% | 5.5% | 9.277% | 35.65% | 5.269% | 3.261% | 7.204%
| PDM3 |  | -0.8093% | 0.5878% | 71.1% | -16.18% | -13.38% | 58.57% | -5.699% | 4929% | 60.66% | -7.795% | 1.985% | -9.192% | 17.35% | -20.37% | -16.18% | -5.699% |  | -20.03% | -3.603% | 46.08% | 143.1% | 210.1% | 32.72% | -5%
| Z-Score (Absolute) |  | 0.04 | 0.04 | 4.3 | NA | 0.8 | 3.54 | 0.34 | NA | NA | 0.46 | 0.13 | 0.55 | 1.05 | 1.22 | NA | 0.34 | 0.3 | 1.2 | 0.21 | 2.79 | 8.64 | 12.69 | 1.98 | 0.3
Results from laboratories 1.07, 1.13 and 1.23 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 234b (Certified Value 0.54 ppm) | [30 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 0.61 | 0.56 | 0.49 | 0.492 | 0.6 | 0.52 | 0.52 | 0.55 | 3.1919 | 0.57 | [19 empty] | 1
| 1 | 2 | 0.61 | 0.57 | 0.52 | 0.51381 | 0.6 | 0.5 | 0.53 | 0.45 | 3.2198 | 0.53 | [19 empty] | 23
| 1 | 3 | 0.62 | 0.56 | 0.49 | 0.51653 | 0.6 | 0.5 | 0.52 | 0.5 | 3.0288 | 0.51 | [19 empty] | 16
| 1 | 4 | 0.62 | 0.55 | 0.52 | 0.50003 | 0.6 | 0.54 | 0.54 | 0.5 | 3.0554 | 0.56 | [19 empty] | 0.54054
| 1 | 5 | 0.6 | 0.57 | 0.53 | 0.50154 | 0.6 | 0.51 | 0.52 | 0.5 | 3.2445 | 0.55 | [19 empty] | 60
| 1 | 6 | 0.61 | 0.54 | 0.51 | 0.51509 | 0.6 | 0.52 | 0.53 | 0.45 | 3.2398 | 0.55
| Mean |  | 0.61167 | 0.55833 | 0.51 | 0.5065 | 0.6 | 0.515 | 0.52667 | 0.49167 | 3.1634 | 0.545
| Median |  | 0.61 | 0.56 | 0.515 | 0.50767 | 0.6 | 0.515 | 0.525 | 0.5 | 3.2058 | 0.55
| Std Dev. |  | 0.0075277 | 0.01169 | 0.016733 | 0.010043 | 0 | 0.015166 | 0.008165 | 0.037639 | 0.096088 | 0.021679
| Rel.Std.Dev. |  | 1.231% | 2.094% | 3.281% | 1.983% | 0% | 2.945% | 1.55% | 7.655% | 3.038% | 3.978%
| PDM3 |  | 13.16% | 3.292% | -5.649% | -6.297% | 11% | -4.724% | -2.566% | -9.041% | 485.2% | 0.8257%
| Z-Score (Absolute) |  | 1.85 | 0.55 | 0.63 | 0.72 | 1.57 | 0.51 | 0.22 | 1.08 | 64.23 | 0.22
| Analytical results for Te in OREAS 234b (Certified Value 0.1 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.16 | < 0.2 | 0.14674 | < 0.1 | 0.2 | < 0.2 | 15 | 0.1 | 0.1 | 0.09 | 0.1 | 0.14 | < 5 | < 0.1 | < 0.5 | 0.07 | < 1 | 0.13432 | 0.11 | 0.1027 | 0.07 | 0.1 | 0.15 | 0.12 | [5 empty] | 1
| 1 | 2 | 0.08 | < 0.2 | 0.14628 | < 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.09 | 0.08 | 0.1 | 0.12 | < 5 | < 0.1 | < 0.5 | 0.08 | < 1 | 0.11578 | 0.11 | 0.0863 | 0.05 | 0.14 | 0.16 | 0.06 | [5 empty] | 24
| 1 | 3 | 0.05 | < 0.2 | 0.15507 | 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.13 | 0.12 | 0.11 | 0.13 | < 5 | 0.1 | < 0.5 | 0.07 | < 1 | 0.12809 | 0.1 | 0.1027 | 0.04 | 0.15 | 0.15 | 0.11 | [5 empty] | 16
| 1 | 4 | 0.07 | < 0.2 | 0.14174 | 0.1 | < 0.1 | < 0.2 | 15 | 0.1 | 0.12 | 0.09 | 0.09 | 0.09 | < 5 | 0.1 | < 0.5 | 0.09 | < 1 | 0.13315 | 0.12 | 0.1356 | 0.05 | 0.12 | 0.12 | 0.08 | [5 empty] | 0.10103
| 1 | 5 | 0.07 | < 0.2 | 0.14875 | 0.1 | < 0.1 | < 0.2 | 16 | < 0.1 | 0.1 | 0.09 | 0.1 | 0.1 | < 5 | 0.1 | < 0.5 | 0.08 | < 1 | 0.13987 | 0.09 | 0.0904 | 0.07 | 0.11 | 0.14 | 0.08 | [5 empty] | 61
| 1 | 6 | 0.13 | < 0.2 | 0.15119 | < 0.1 | 0.1 | < 0.2 | 15 | < 0.1 | 0.11 | 0.1 | 0.1 | 0.16 | < 5 | < 0.1 | < 0.5 | 0.09 | < 1 | 0.1247 | 0.1 | 0.0827 | 0.09 | 0.1 | 0.15 | 0.11
| Mean |  | 0.093333 |  | 0.1483 | 0.1 | 0.15 |  | 15.1667 | 0.1 | 0.10833 | 0.095 | 0.1 | 0.12333 |  | 0.1 |  | 0.08 |  | 0.12932 | 0.105 | 0.10007 | 0.061667 | 0.12 | 0.145 | 0.093333
| Median |  | 0.075 |  | 0.14775 | 0.1 | 0.15 |  | 15 | 0.1 | 0.105 | 0.09 | 0.1 | 0.125 |  | 0.1 |  | 0.08 |  | 0.13062 | 0.105 | 0.09655 | 0.06 | 0.115 | 0.15 | 0.095
| Std Dev. |  | 0.042269 |  | 0.0045599 | 1.6997e-17 | 0.070711 |  | 0.40825 | 0 | 0.01472 | 0.013784 | 0.0063246 | 0.02582 |  | 1.6997e-17 |  | 0.0089443 |  | 0.0084517 | 0.010488 | 0.019293 | 0.018348 | 0.020976 | 0.013784 | 0.023381
| Rel.Std.Dev. |  | 45.29% |  | 3.075% | 1.7e-14% | 47.14% |  | 2.692% | 0% | 13.59% | 14.51% | 6.325% | 20.94% |  | 1.7e-14% |  | 11.18% |  | 6.536% | 9.989% | 19.28% | 29.75% | 17.48% | 9.506% | 25.05%
| PDM3 |  | -7.619% |  | 46.78% | -1.02% | 48.47% |  | 14912% | -1.02% | 7.228% | -5.969% | -1.02% | 22.08% |  | -1.02% |  | -20.82% |  | 28% | 3.929% | -0.9543% | -38.96% | 18.78% | 43.52% | -7.619%
| Z-Score (Absolute) |  | 0.22 | 0 | 1.63 | 0.84 | 0.56 | 0 | NA | 1.12 | 0.28 | 0.17 | 0 | 0.79 | 80.92 | 0.84 | 5.06 | 0.67 | 13.49 | 0.99 | 0.17 | 0 | 1.29 | 0.67 | 1.52 | 0.22
Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 234b (Certified Value 1.02 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1 | 1.03 | 1.03 | 0.84281 | 0.97974 | 1.1 | 1.12 | 1.15 | 0.96 | 1 | 1 | 1.05 | 0.99 | 1 | 0.85 | 0.9 | 1.05 | 1.1 | 1.0722 | 1.01 | 1.2452 | 1.7 | 1.1 | 1.1 | 0.9 |  |  |  |  | 1
| 1 | 2 | 1.5 | 1.04 | 0.95 | 0.82886 | 0.99224 | 1.1 | 1.07 | 1.06 | 1.04 | 1 | 0.93 | 1.03 | 0.95 | 1.02 | 0.96 | 0.9 | 1.07 | 1 | 1.0593 | 1.03 | 1.2365 | 1.5 | 1 | 1.1 | 0.9 |  |  |  |  | 25
| 1 | 3 | 2.1 | 1.02 | 0.87 | 0.83618 | 1.0088 | 1.2 | 1.07 | 1.15 | 1.01 | 1 | 0.98 | 1.01 | 0.98 | 1 | 0.83 | 0.9 | 1.07 | 1 | 0.98341 | 1.03 | 1.2735 | 1.5 | 1 | 1.1 | 0.9 |  |  |  |  | 16
| 1 | 4 | 2.3 | 0.99 | 0.93 | 0.83875 | 1.0132 | 1.1 | 1.1 | 1.16 | 0.98 | 1 | 0.94 | 1.04 | 0.97 | 0.95 | 0.87 | 0.9 | 1.08 | 1.1 | 0.98211 | 1.01 | 1.2169 | 1.5 | 1.1 | 1.1 | 1.1 |  |  |  |  | 1.0159
| 1 | 5 | 2 | 1.02 | 0.97 | 0.82223 | 0.96733 | 1 | 1.1 | 1.13 | 0.98 | 1 | 0.94 | 1.08 | 1.06 | 0.99 | 0.93 | 0.9 | 1.03 | 1.1 | 1.0718 | 0.93 | 1.2348 | 1.3 | 1 | 1.1 | 1 |  |  |  |  | 62
| 1 | 6 | 2.6 | 1 | 0.93 | 0.83802 | 0.9941 | 1 | 1.08 | 1.11 | 1.02 | 1 | 0.93 | 1.08 | 1.01 | 1.02 | 0.89 | 0.9 | 1.08 | 1.1 | 1.0722 | 0.97 | 1.2407 | 1.4 | 1 | 1.3 | 1.1
| Mean |  | 1.9167 | 1.0167 | 0.94667 | 0.83448 | 0.99256 | 1.0833 | 1.09 | 1.1267 | 0.99833 | 1 | 0.95333 | 1.0483 | 0.99333 | 0.99667 | 0.88833 | 0.9 | 1.0633 | 1.0667 | 1.0402 | 0.99667 | 1.2413 | 1.4833 | 1.0333 | 1.1333 | 0.98333
| Median |  | 2.05 | 1.02 | 0.94 | 0.8371 | 0.99317 | 1.1 | 1.09 | 1.14 | 0.995 | 1 | 0.94 | 1.045 | 0.985 | 1 | 0.88 | 0.9 | 1.07 | 1.1 | 1.0655 | 1.01 | 1.2386 | 1.5 | 1 | 1.1 | 0.95
| Std Dev. |  | 0.57764 | 0.018619 | 0.052789 | 0.0075431 | 0.017282 | 0.075277 | 0.02 | 0.037238 | 0.029944 | 0 | 0.029439 | 0.027869 | 0.038297 | 0.02582 | 0.04916 | 0 | 0.019664 | 0.05164 | 0.04476 | 0.039328 | 0.01851 | 0.13292 | 0.05164 | 0.08165 | 0.098319
| Rel.Std.Dev. |  | 30.14% | 1.831% | 5.576% | 0.9039% | 1.741% | 6.949% | 1.835% | 3.305% | 2.999% | 0% | 3.088% | 2.658% | 3.855% | 2.591% | 5.534% | 0% | 1.849% | 4.841% | 4.303% | 3.946% | 1.491% | 8.961% | 4.997% | 7.204% | 9.999%
| PDM3 |  | 88.67% | 0.07671% | -6.814% | -17.86% | -2.296% | 6.639% | 7.295% | 10.9% | -1.728% | -1.564% | -6.158% | 3.194% | -2.22% | -1.892% | -12.56% | -11.41% | 4.67% | 4.999% | 2.392% | -1.892% | 22.19% | 46.01% | 1.717% | 11.56% | -3.204%
| Z-Score (Absolute) |  | 12.03 | 0.11 | 1.06 | 2.57 | 0.44 | 0.79 | 0.88 | 1.37 | 0.36 | NA | 0.97 | 0.31 | 0.43 | 0.38 | 1.84 | 1.69 | 0.52 | 0.56 | 0.2 | 0.38 | 2.92 | 6.18 | 0.11 | 1.46 | 0.56
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 234b (Certified Value 0.592 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.59 | 0.601 | 0.58673 | 0.57941 | 0.517 | 0.6414 | 0.598 | 0.5754 | 0.57 | 0.59 | 0.628 | 0.601 | 0.619 | 0.585 | 0.57 | 0.6048 | 0.592 | 0.612 | 0.82 | 0.57946 | 0.582 | 0.58116 | 0.591 | 0.51769 | 0.59 | 0.61 | 0.59 | 0.7 |  | 1
| 1 | 2 | 0.6 | 0.5969 | 0.58883 | 0.61039 | 0.569 | 0.6474 | 0.59 | 0.6077 | 0.55 | 0.597 | 0.603 | 0.584 | 0.572 | 0.595 | 0.57 | 0.6193 | 0.59 | 0.588 | 0.66 | 0.5628 | 0.58 | 0.575 | 0.582 | 0.51709 | 0.57 | 0.58 | 0.59 | 0.7 |  | 26
| 1 | 3 | 0.6 | 0.6104 | 0.58412 | 0.62039 | 0.602 | 0.6295 | 0.595 | 0.5924 | 0.56 | 0.615 | 0.621 | 0.57 | 0.593 | 0.595 | 0.57 | 0.6114 | 0.592 | 0.583 | 0.63 | 0.54734 | 0.579 | 0.6019 | 0.589 | 0.51399 | 0.57 | 0.58 | 0.59 | 0.7 |  | 16
| 1 | 4 | 0.59 | 0.6057 | 0.58061 | 0.60459 | 0.594 | 0.6414 | 0.588 | 0.6015 | 0.56 | 0.604 | 0.606 | 0.565 | 0.57 | 0.588 | 0.54 | 0.606 | 0.587 | 0.585 | 0.64 | 0.5648 | 0.576 | 0.62957 | 0.59 | 0.50569 | 0.58 | 0.59 | 0.59 | 0.63 |  | 0.59177
| 1 | 5 | 0.59 | 0.5896 | 0.58883 | 0.56063 | 0.595 | 0.6355 | 0.601 | 0.6058 | 0.57 | 0.605 | 0.595 | 0.573 | 0.597 | 0.596 | 0.57 | 0.6221 | 0.583 | 0.616 | 0.66 | 0.56393 | 0.579 | 0.5939 | 0.593 | 0.52069 | 0.59 | 0.6 | 0.59 | 0.65 |  | 63
| 1 | 6 | 0.61 | 0.5716 | 0.58802 | 0.60333 | 0.248 | 0.6355 | 0.587 | 0.5836 | 0.56 | 0.597 | 0.615 | 0.582 | 0.596 | 0.595 | 0.53 | 0.6054 | 0.587 | 0.601 | 0.645 | 0.55103 | 0.579 | 0.58908 | 0.563 | 0.51899 | 0.57 | 0.59 | 0.59 | 0.63
| Mean |  | 0.59667 | 0.59587 | 0.58619 | 0.59646 | 0.52083 | 0.63845 | 0.59317 | 0.5944 | 0.56167 | 0.60133 | 0.61133 | 0.57917 | 0.59117 | 0.59233 | 0.55833 | 0.6115 | 0.5885 | 0.5975 | 0.67583 | 0.56156 | 0.57917 | 0.5951 | 0.58467 | 0.51569 | 0.57833 | 0.59167 | 0.59 | 0.66833
| Median |  | 0.595 | 0.59895 | 0.58738 | 0.60396 | 0.5815 | 0.63845 | 0.5925 | 0.59695 | 0.56 | 0.6005 | 0.6105 | 0.5775 | 0.5945 | 0.595 | 0.57 | 0.6087 | 0.5885 | 0.5945 | 0.6525 | 0.56337 | 0.579 | 0.59149 | 0.5895 | 0.51739 | 0.575 | 0.59 | 0.59 | 0.675
| Std Dev. |  | 0.008165 | 0.013881 | 0.0032525 | 0.022155 | 0.13727 | 0.0062452 | 0.0057067 | 0.012953 | 0.0075277 | 0.008641 | 0.012242 | 0.012891 | 0.01817 | 0.0046332 | 0.018348 | 0.0075572 | 0.0035071 | 0.014293 | 0.071583 | 0.011412 | 0.0019408 | 0.019342 | 0.011255 | 0.0053803 | 0.0098319 | 0.01169 | 0 | 0.035449
| Rel.Std.Dev. |  | 1.368% | 2.33% | 0.5549% | 3.714% | 26.36% | 0.9782% | 0.9621% | 2.179% | 1.34% | 1.437% | 2.003% | 2.226% | 3.074% | 0.7822% | 3.286% | 1.236% | 0.5959% | 2.392% | 10.59% | 2.032% | 0.3351% | 3.25% | 1.925% | 1.043% | 1.7% | 1.976% | 0% | 5.304%
| PDM3 |  | 0.8282% | 0.693% | -0.9424% | 0.793% | -11.99% | 7.889% | 0.2367% | 0.4451% | -5.086% | 1.617% | 3.307% | -2.129% | -0.1013% | 0.09589% | -5.65% | 3.335% | -0.5519% | 0.969% | 14.21% | -5.104% | -2.129% | 0.5637% | -1.2% | -12.86% | -2.27% | -0.01677% | -0.2984% | 12.94%
| Z-Score (Absolute) |  | 0.38 | 0.31 | 0.47 | 0.36 | 5.76 | 3.76 | 0.09 | 0.19 | 2.45 | 0.76 | 1.56 | 1.04 | 0.07 | 0.03 | 2.72 | 1.58 | 0.28 | 0.45 | 6.78 | 2.46 | 1.04 | 0.25 | 0.59 | 6.17 | 1.11 | 0.03 | 0.16 | 6.18
| Analytical results for Tl in OREAS 234b (Certified Value 0.23 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.21 | 0.24 | 0.22577 | 0.22 | 0.24 | 0.23 | 0.1 | 0.22 | 0.25 | 0.22 | 0.24 | < 5 | 0.2 | < 0.5 | 0.22 | 1.28 | 0.2 | 0.24555 | 0.24 | 0.2435 | 0.19 | 0.23 | 0.24 | 0.21 | [5 empty] | 1
| 1 | 2 | 0.21 | 0.24 | 0.24027 | 0.23 | 0.2 | 0.24 | 0.1 | 0.22 | 0.24 | 0.21 | 0.24 | < 5 | 0.24 | < 0.5 | 0.23 | 0.58 | 0.2 | 0.23624 | 0.23 | 0.2399 | 0.19 | 0.22 | 0.25 | 0.22 | [5 empty] | 27
| 1 | 3 | 0.21 | 0.25 | 0.23575 | 0.23 | 0.25 | 0.23 | < 0.1 | 0.23 | 0.23 | 0.23 | 0.23 | < 5 | 0.19 | < 0.5 | 0.22 | 0.98 | 0.2 | 0.22218 | 0.23 | 0.2456 | 0.18 | 0.21 | 0.23 | 0.22 | [5 empty] | 16
| 1 | 4 | 0.21 | 0.23 | 0.23676 | 0.22 | 0.24 | 0.24 | 0.1 | 0.2 | 0.23 | 0.22 | 0.22 | < 5 | 0.21 | < 0.5 | 0.22 | < 0.1 | 0.2 | 0.22943 | 0.23 | 0.2364 | 0.18 | 0.23 | 0.22 | 0.23 | [5 empty] | 0.2266
| 1 | 5 | 0.2 | 0.24 | 0.22197 | 0.21 | 0.23 | 0.23 | < 0.1 | 0.22 | 0.25 | 0.24 | 0.23 | < 5 | 0.22 | < 0.5 | 0.22 | 0.46 | 0.2 | 0.23801 | 0.2 | 0.2365 | 0.19 | 0.22 | 0.24 | 0.22 | [5 empty] | 64
| 1 | 6 | 0.21 | 0.22 | 0.25016 | 0.22 | 0.22 | 0.24 | 0.1 | 0.21 | 0.25 | 0.22 | 0.23 | < 5 | 0.22 | < 0.5 | 0.22 | 0.39 | 0.25 | 0.24174 | 0.23 | 0.2349 | 0.18 | 0.22 | 0.25 | 0.24
| Mean |  | 0.20833 | 0.23667 | 0.23511 | 0.22167 | 0.23 | 0.235 | 0.1 | 0.21667 | 0.24167 | 0.22333 | 0.23167 |  | 0.21333 |  | 0.22167 | 0.738 | 0.20833 | 0.23553 | 0.22667 | 0.23947 | 0.185 | 0.22167 | 0.23833 | 0.22333
| Median |  | 0.21 | 0.24 | 0.23626 | 0.22 | 0.235 | 0.235 | 0.1 | 0.22 | 0.245 | 0.22 | 0.23 |  | 0.215 |  | 0.22 | 0.58 | 0.2 | 0.23713 | 0.23 | 0.2382 | 0.185 | 0.22 | 0.24 | 0.22
| Std Dev. |  | 0.0040825 | 0.010328 | 0.010162 | 0.0075277 | 0.017889 | 0.0054772 | 0 | 0.010328 | 0.0098319 | 0.010328 | 0.0075277 |  | 0.017512 |  | 0.0040825 | 0.37937 | 0.020412 | 0.0084977 | 0.013663 | 0.0043159 | 0.0054772 | 0.0075277 | 0.01169 | 0.010328
| Rel.Std.Dev. |  | 1.96% | 4.364% | 4.322% | 3.396% | 7.778% | 2.331% | 0% | 4.767% | 4.068% | 4.624% | 3.249% |  | 8.209% |  | 1.842% | 51.4% | 9.798% | 3.608% | 6.028% | 1.802% | 2.961% | 3.396% | 4.905% | 4.624%
| PDM3 |  | -8.062% | 4.442% | 3.756% | -2.178% | 1.5% | 3.706% | -55.87% | -4.385% | 6.648% | -1.442% | 2.235% |  | -5.856% |  | -2.178% | 225.7% | -8.062% | 3.938% | 0.02851% | 5.677% | -18.36% | -2.178% | 5.177% | -1.442%
| Z-Score (Absolute) |  | 1.75 | 0.54 | 0.41 | 0.67 | 0 | 0.4 | NA | 1.08 | 0.94 | 0.54 | 0.13 | 183.68 | 1.35 | 1.62 | 0.67 | 31.83 | 1.75 | 0.45 | 0.27 | 0.77 | 3.64 | 0.67 | 0.67 | 0.54
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 234b (Certified Value 0.31 ppm) | [30 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | [22 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [22 empty] | 2
3
| 1 | 1 | 0.33 | 0.28 | 0.28155 | 0.3 | 0.3 | 0.35 | 0.34016 | [22 empty] | 1
| 1 | 2 | 0.33 | 0.3 | 0.28401 | 0.3 | 0.31 | 0.35 | 0.34803 | [22 empty] | 28
| 1 | 3 | 0.32 | 0.28 | 0.28387 | 0.3 | 0.32 | 0.3 | 0.33016 | [22 empty] | 16
| 1 | 4 | 0.32 | 0.29 | 0.28959 | 0.3 | 0.31 | 0.35 | 0.31691 | [22 empty] | 0.31326
| 1 | 5 | 0.33 | 0.31 | 0.29082 | 0.3 | 0.29 | 0.35 | 0.33837 | [22 empty] | 65
| 1 | 6 | 0.36 | 0.29 | 0.28469 | 0.3 | 0.31 | 0.3 | 0.34337
| Mean |  | 0.33167 | 0.29167 | 0.28576 | 0.3 | 0.30667 | 0.33333 | 0.33617
| Median |  | 0.33 | 0.29 | 0.28435 | 0.3 | 0.31 | 0.35 | 0.33927
| Std Dev. |  | 0.01472 | 0.01169 | 0.0036259 | 0 | 0.010328 | 0.02582 | 0.011138
| Rel.Std.Dev. |  | 4.438% | 4.008% | 1.269% | 0% | 3.368% | 7.746% | 3.313%
| PDM3 |  | 5.874% | -6.895% | -8.781% | -4.234% | -2.106% | 6.406% | 7.311%
| Z-Score (Absolute) |  | 0.54 | 1.19 | 1.45 | NA | 0.54 | 0.61 | 0.74
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 234b (Certified Value 0.3 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.4 | 0.31 | 0.32 | 0.2399 | 0.32897 | 0.3 | 0.4 | 0.33 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.24 | 0.3 | 0.32 | 0.4 | 0.31215 | 0.3 | 0.3261 | 0.313 | 0.3 | 0.3 | 0.281 |  |  |  | 1
| 1 | 2 | 0.4 | 0.33 | 0.31 | 0.24347 | 0.34165 | 0.3 | 0.38 | 0.3 | 0.32 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.28 | 0.3 | 0.32 | 0.3 | 0.30122 | 0.3 | 0.3291 | 0.308 | 0.3 | 0.3 | 0.279 |  |  |  | 29
| 1 | 3 | 0.5 | 0.33 | 0.31 | 0.23632 | 0.32184 | 0.3 | 0.37 | 0.3 | 0.35 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.23 | 0.3 | 0.32 | 0.3 | 0.29151 | 0.3 | 0.3301 | 0.303 | 0.3 | 0.3 | 0.292 |  |  |  | 16
| 1 | 4 | 0.4 | 0.32 | 0.32 | 0.24032 | 0.33211 | 0.3 | 0.39 | 0.3 | 0.32 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.25 | 0.3 | 0.34 | 0.3 | 0.29124 | 0.3 | 0.3236 | 0.311 | 0.3 | 0.3 | 0.326 |  |  |  | 0.30383
| 1 | 5 | 0.4 | 0.33 | 0.33 | 0.24431 | 0.344 | 0.3 | 0.38 | 0.31 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.27 | 0.3 | 0.32 | 0.3 | 0.30581 | 0.3 | 0.3264 | 0.315 | 0.3 | 0.3 | 0.302 |  |  |  | 66
| 1 | 6 | 0.5 | 0.31 | 0.33 | 0.22931 | 0.32748 | 0.3 | 0.37 | 0.29 | 0.32 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | < 10 | 0.25 | 0.3 | 0.32 | 0.3 | 0.31116 | 0.3 | 0.3251 | 0.307 | 0.3 | 0.3 | 0.342
| Mean |  | 0.43333 | 0.32167 | 0.32 | 0.23894 | 0.33267 | 0.3 | 0.38167 | 0.305 | 0.31833 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 |  | 0.25333 | 0.3 | 0.32333 | 0.31667 | 0.30218 | 0.3 | 0.32673 | 0.3095 | 0.3 | 0.3 | 0.30367
| Median |  | 0.4 | 0.325 | 0.32 | 0.24011 | 0.33054 | 0.3 | 0.38 | 0.3 | 0.32 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 |  | 0.25 | 0.3 | 0.32 | 0.3 | 0.30351 | 0.3 | 0.32625 | 0.3095 | 0.3 | 0.3 | 0.297
| Std Dev. |  | 0.05164 | 0.0098319 | 0.0089443 | 0.0055103 | 0.0085692 | 0 | 0.01169 | 0.013784 | 0.018348 | 0 | 0 | 0 | 0 | 0 |  | 0.018619 | 0 | 0.008165 | 0.040825 | 0.0092485 | 0 | 0.0024468 | 0.0043704 | 0 | 0 | 0.025414
| Rel.Std.Dev. |  | 11.92% | 3.057% | 2.795% | 2.306% | 2.576% | 0% | 3.063% | 4.519% | 5.764% | 0% | 0% | 0% | 0% | 0% |  | 7.35% | 0% | 2.525% | 12.89% | 3.061% | 0% | 0.7489% | 1.412% | 0% | 0% | 8.369%
| PDM3 |  | 42.62% | 5.869% | 5.321% | -21.36% | 9.492% | -1.262% | 25.62% | 0.3837% | 4.772% | -1.262% | -1.262% | -1.262% | -1.262% | -1.262% |  | -16.62% | -1.262% | 6.418% | 4.224% | -0.5439% | -1.262% | 7.537% | 1.865% | -1.262% | -1.262% | -0.05513%
| Z-Score (Absolute) |  | 18.51 | 2.66 | 2.42 | 9.08 | 4.22 | 0.42 | 11.18 | 0.29 | 2.19 | 0.42 | 0.42 | 0.42 | 0.42 | 0.42 | 666.8 | 7.04 | 0.42 | 2.9 | 1.95 | 0.11 | 0.42 | 3.38 | 0.93 | 0.42 | 0.42 | 0.11
| Analytical results for V in OREAS 234b (Certified Value 258 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 280 | 261 | 233.35 | 250.6502 | 243.9626 | 239 | 321 | 248 | 246 | 251 | 254 | 272 | 265 | 273 | 257 | 269 | 258 | 261 | 274 | 264.92 | 266.7632 | 255 | 263.2656 | 258 | 241.983 | 269 | 263 | 272 | 218 | 1
| 1 | 2 | 279 | 265 | 233.16 | 248.7063 | 258.3093 | 263 | 323 | 255 | 268 | 252 | 258 | 263 | 251 | 255 | 254 | 268 | 266 | 262 | 258 | 221.39 | 257.3608 | 255 | 263.7454 | 253 | 239.9528 | 263 | 252 | 272 | 227 | 30
| 1 | 3 | 269 | 275 | 230.06 | 247.5601 | 263.4129 | 279 | 318 | 251 | 263 | 251 | 254 | 270 | 249 | 263 | 258 | 269 | 264 | 265 | 263 | 220.99 | 247.4248 | 255 | 265.8789 | 251 | 242.2521 | 259 | 253 | 275 | 236 | 16
| 1 | 4 | 264 | 266 | 229.7 | 248.2278 | 255.4179 | 257 | 318 | 250 | 259 | 249 | 255 | 262 | 246 | 253 | 254 | 252 | 266 | 261 | 263 | 211.12 | 255.2448 | 255 | 274.8136 | 256 | 235.0524 | 268 | 257 | 273 | 258 | 257.6248
| 1 | 5 | 260 | 254 | 232.82 | 243.981 | 238.6903 | 252 | 317 | 254 | 243 | 249 | 252 | 259 | 248 | 264 | 260 | 267 | 268 | 257 | 274 | 224.55 | 256.0544 | 255 | 267.5656 | 257 | 242.4692 | 266 | 255 | 271 | 242 | 67
| 1 | 6 | 269 | 248 | 232.49 | 249.1468 | 253.5969 | 131 | 321 | 251 | 255 | 248 | 252 | 264 | 256 | 263 | 260 | 248 | 258 | 262 | 267 | 213.72 | 246.1644 | 255 | 267.8358 | 245 | 241.3842 | 263 | 254 | 272 | 274
| Mean |  | 270.1667 | 261.5 | 231.93 | 248.0454 | 252.2317 | 236.8333 | 319.6667 | 251.5 | 255.6667 | 250 | 254.1667 | 265 | 252.5 | 261.8333 | 257.1667 | 262.1667 | 263.3333 | 261.3333 | 266.5 | 226.115 | 254.8354 | 255 | 267.1841 | 253.3333 | 240.5156 | 264.6667 | 255.6667 | 272.5 | 242.5
| Median |  | 269 | 263 | 232.655 | 248.467 | 254.5074 | 254.5 | 319.5 | 251 | 257 | 250 | 254 | 263.5 | 250 | 263 | 257.5 | 267.5 | 265 | 261.5 | 265 | 221.19 | 255.6496 | 255 | 266.7223 | 254.5 | 241.6836 | 264.5 | 254.5 | 272 | 239
| Std Dev. |  | 7.9854 | 9.5237 | 1.6189 | 2.246 | 9.2273 | 53.488 | 2.3381 | 2.5884 | 9.7091 | 1.5492 | 2.2286 | 4.98 | 7.0071 | 7.1671 | 2.7142 | 9.5376 | 4.3205 | 2.582 | 6.473 | 19.6741 | 7.4887 | 0 | 4.1872 | 4.8442 | 2.825 | 3.7238 | 3.9833 | 1.3784 | 20.5694
| Rel.Std.Dev. |  | 2.956% | 3.642% | 0.698% | 0.9055% | 3.658% | 22.58% | 0.7314% | 1.029% | 3.798% | 0.6197% | 0.8768% | 1.879% | 2.775% | 2.737% | 1.055% | 3.638% | 1.641% | 0.988% | 2.429% | 8.701% | 2.939% | 0% | 1.567% | 1.912% | 1.175% | 1.407% | 1.558% | 0.5058% | 8.482%
| PDM3 |  | 4.868% | 1.504% | -9.974% | -3.718% | -2.093% | -8.07% | 24.08% | -2.377% | -0.7601% | -2.96% | -1.342% | 2.863% | -1.989% | 1.634% | -0.1778% | 1.763% | 2.216% | 1.439% | 3.445% | -12.23% | -1.083% | -1.019% | 3.711% | -1.666% | -6.641% | 2.733% | -0.7601% | 5.774% | -5.871%
| Z-Score (Absolute) |  | 1.5 | 0.61 | 2.46 | 0.79 | 0.36 | 1.95 | 6.64 | 0.43 | 0 | 0.59 | 0.16 | 0.97 | 0.33 | 0.64 | 0.16 | 0.67 | 0.8 | 0.59 | 1.12 | 3.07 | 0.09 | 0.07 | 1.19 | 0.24 | 1.57 | 0.93 | 0 | 1.75 | 1.37
| Analytical results for W in OREAS 234b (Certified Value 33.3 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 36.7 | 32.9 | 35.24 | 25.411 | 31.9167 | 7.5 | 33.3 | 31.6 | 31.9 | 31 | 6.5 | 37 | 35.5 | 30.3 | 34.8 | 29 | 26.9 | 30.2 | 34 | 37 | 32.8653 | 35.2 | 35.8115 | 35.5 | 32.3 | 36 | 31.9 |  |  | 1
| 1 | 2 | 37.5 | 33.1 | 37.02 | 25.3944 | 35.2135 | 22.3 | 31.7 | 30.8 | 32.7 | 33 | 6.2 | 35 | 34.9 | 29.9 | 34 | 29 | 29.2 | 30.2 | 32 | 37 | 32.7564 | 34.3 | 35.6797 | 34.9 | 31.8 | 37 | 31.1 |  |  | 31
| 1 | 3 | 35.9 | 33.4 | 34.57 | 25.6506 | 34.5424 | 30.8 | 32.8 | 32 | 31.9 | 32 | 5.6 | 35.2 | 34.6 | 30 | 32.8 | 29 | 27.5 | 30.4 | 32 | 36 | 31.9271 | 35.5 | 36.1449 | 35.3 | 30.8 | 35 | 32.2 |  |  | 16
| 1 | 4 | 36.1 | 32.4 | 34.46 | 25.4201 | 35.1492 | 28.2 | 33 | 31.5 | 33.4 | 32 | 5.6 | 35.6 | 35.4 | 29.7 | 33.5 | 26 | 27.7 | 31.8 | 33 | 37 | 31.3874 | 34.2 | 35.2661 | 35.7 | 32.9 | 36 | 35.6 |  |  | 33.305
| 1 | 5 | 35.3 | 33.5 | 35.71 | 25.6179 | 32.3838 | 26.2 | 33.1 | 31.8 | 32.4 | 33 | 5 | 34.7 | 34.7 | 31.8 | 34.8 | 29 | 28.7 | 29.8 | 35 | 39 | 33.5127 | 32.2 | 35.3689 | 36 | 32.5 | 35.9 | 34.2 |  |  | 68
| 1 | 6 | 36.1 | 32 | 36.59 | 25.5262 | 33.5654 | 0.8 | 33.4 | 31 | 32.2 | 33 | 5.2 | 34.7 | 34.3 | 32.1 | 36.7 | 26 | 28 | 30.8 | 34 | 39 | 33.4599 | 32.7 | 35.3447 | 34.9 | 32.1 | 36.3 | 36.5
| Mean |  | 36.2667 | 32.8833 | 35.5983 | 25.5034 | 33.7952 | 19.3 | 32.8833 | 31.45 | 32.4167 | 32.3333 | 5.6833 | 35.3667 | 34.9 | 30.6333 | 34.4333 | 28 | 28 | 30.5333 | 33.3333 | 37.5 | 32.6515 | 34.0167 | 35.6026 | 35.3833 | 32.0667 | 36.0333 | 33.5833
| Median |  | 36.1 | 33 | 35.475 | 25.4731 | 34.0539 | 24.25 | 33.05 | 31.55 | 32.3 | 32.5 | 5.6 | 35.1 | 34.8 | 30.15 | 34.4 | 29 | 27.85 | 30.3 | 33.5 | 37 | 32.8108 | 34.25 | 35.5243 | 35.4 | 32.2 | 36 | 33.2
| Std Dev. |  | 0.75277 | 0.58452 | 1.0487 | 0.11194 | 1.4126 | 12.2439 | 0.61779 | 0.46368 | 0.57067 | 0.8165 | 0.57417 | 0.86872 | 0.46904 | 1.0424 | 1.3515 | 1.5492 | 0.83427 | 0.70048 | 1.2111 | 1.2247 | 0.84554 | 1.3227 | 0.34001 | 0.44008 | 0.72296 | 0.64704 | 2.1849
| Rel.Std.Dev. |  | 2.076% | 1.778% | 2.946% | 0.4389% | 4.18% | 63.44% | 1.879% | 1.474% | 1.76% | 2.525% | 10.1% | 2.456% | 1.344% | 3.403% | 3.925% | 5.533% | 2.98% | 2.294% | 3.633% | 3.266% | 2.59% | 3.889% | 0.955% | 1.244% | 2.255% | 1.796% | 6.506%
| PDM3 |  | 8.892% | -1.266% | 6.886% | -23.42% | 1.472% | -42.05% | -1.266% | -5.57% | -2.667% | -2.918% | -82.94% | 6.19% | 4.789% | -8.022% | 3.388% | -15.93% | -15.93% | -8.322% | 0.08495% | 12.6% | -1.962% | 2.137% | 6.899% | 6.24% | -3.718% | 8.192% | 0.8356%
| Z-Score (Absolute) |  | 1.01 | 0 | 0.81 | 2.21 | 0.27 | 4.07 | 0 | 0.43 | 0.14 | 0.16 | 8.15 | 0.74 | 0.6 | 0.67 | 0.46 | 1.46 | 1.46 | 0.7 | 0.13 | 1.38 | 0.07 | 0.34 | 0.81 | 0.75 | 0.24 | 0.94 | 0.21
| Analytical results for Y in OREAS 234b (Certified Value 19.4 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 18.4 | 19.36 | 17.12 | 17.9596 | 18.0993 | 18.7 | 20 | 20.6 | 18.64 | 18 | 20.2 | 21.1 | 20.8 | 19.8 | 21 | 19 | 17.3 | 18.7 | 19.1 | 20.6 | 19.8601 | 20.2 | 19.9359 | 20.3 | 20.1 | 21.5 | 17.1 |  |  | 1
| 1 | 2 | 18.5 | 19.36 | 18.02 | 17.9235 | 19.1621 | 19.8 | 18.8 | 20.2 | 19.02 | 18 | 20.1 | 20.3 | 20.5 | 19.4 | 20.7 | 19 | 18.4 | 19.4 | 18.6 | 19.7 | 20.0502 | 20.3 | 20.1429 | 20 | 18.5 | 21.9 | 16.9 |  |  | 32
| 1 | 3 | 18.7 | 19.38 | 16.93 | 17.92 | 19.6202 | 20.1 | 19.2 | 20.8 | 18.73 | 18 | 19.7 | 20.4 | 20.5 | 20.6 | 20.7 | 19 | 16.8 | 18.7 | 19.7 | 20 | 20.0723 | 20.3 | 20.0174 | 20.1 | 19 | 21.2 | 17.5 |  |  | 16
| 1 | 4 | 18.6 | 19.47 | 17.68 | 17.8936 | 19.3051 | 18.2 | 19.6 | 19.9 | 19.4 | 18 | 20.6 | 20.1 | 21.3 | 19.6 | 20.9 | 18 | 16.4 | 18.5 | 19.2 | 20.8 | 20.2702 | 20.3 | 19.4346 | 20.3 | 19.5 | 20.9 | 19.5 |  |  | 19.4404
| 1 | 5 | 18.5 | 19.38 | 17.99 | 18.0226 | 17.8536 | 18.3 | 19.7 | 21 | 19.1 | 19 | 20.5 | 19.9 | 20.6 | 21 | 21.6 | 19 | 16.8 | 18.2 | 18.9 | 21.3 | 20.1976 | 20.2 | 19.6624 | 20.5 | 19.1 | 21 | 18.1 |  |  | 69
| 1 | 6 | 19.1 | 18.35 | 17.83 | 18.06 | 19.0462 | 18.4 | 19.7 | 20.2 | 18.51 | 18 | 20.3 | 20.1 | 21.2 | 21.9 | 21.9 | 18 | 18 | 18.4 | 20.6 | 20.6 | 19.9579 | 19 | 19.4428 | 20.4 | 19.8 | 21.7 | 20.6
| Mean |  | 18.6333 | 19.2167 | 17.595 | 17.9632 | 18.8477 | 18.9167 | 19.5 | 20.45 | 18.9 | 18.1667 | 20.2333 | 20.3167 | 20.8167 | 20.3833 | 21.1333 | 18.6667 | 17.2833 | 18.65 | 19.35 | 20.5 | 20.0681 | 20.05 | 19.7727 | 20.2667 | 19.3333 | 21.3667 | 18.2833
| Median |  | 18.55 | 19.37 | 17.755 | 17.9416 | 19.1042 | 18.55 | 19.65 | 20.4 | 18.875 | 18 | 20.25 | 20.2 | 20.7 | 20.2 | 20.95 | 19 | 17.05 | 18.6 | 19.15 | 20.6 | 20.0612 | 20.25 | 19.7992 | 20.3 | 19.3 | 21.35 | 17.8
| Std Dev. |  | 0.25033 | 0.42655 | 0.46194 | 0.06511 | 0.70605 | 0.82321 | 0.42895 | 0.41833 | 0.33257 | 0.40825 | 0.32042 | 0.42151 | 0.35449 | 0.96419 | 0.50067 | 0.5164 | 0.77567 | 0.41352 | 0.71204 | 0.57271 | 0.15057 | 0.51672 | 0.30294 | 0.18619 | 0.58195 | 0.39833 | 1.4703
| Rel.Std.Dev. |  | 1.343% | 2.22% | 2.625% | 0.3625% | 3.746% | 4.352% | 2.2% | 2.046% | 1.76% | 2.247% | 1.584% | 2.075% | 1.703% | 4.73% | 2.369% | 2.766% | 4.488% | 2.217% | 3.68% | 2.794% | 0.7503% | 2.577% | 1.532% | 0.9187% | 3.01% | 1.864% | 8.042%
| PDM3 |  | -4.151% | -1.151% | -9.492% | -7.598% | -3.048% | -2.694% | 0.3067% | 5.193% | -2.78% | -6.552% | 4.079% | 4.508% | 7.08% | 4.851% | 8.708% | -3.98% | -11.1% | -4.066% | -0.4649% | 5.451% | 3.229% | 3.136% | 1.709% | 4.25% | -0.5506% | 9.909% | -5.952%
| Z-Score (Absolute) |  | 0.55 | 0.1 | 1.34 | 1.06 | 0.38 | 0.33 | 0.11 | 0.84 | 0.34 | 0.9 | 0.67 | 0.74 | 1.12 | 0.79 | 1.36 | 0.52 | 1.58 | 0.53 | 0 | 0.88 | 0.55 | 0.53 | 0.32 | 0.7 | 0.01 | 1.54 | 0.81
| Analytical results for Yb in OREAS 234b (Certified Value 2.03 ppm) | [30 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 2.3 | 2.06 | 1.69 | 1.9266 | 2 | 2.03 | 2.1 | 1.7 | 2.1968 | 2.1 | [19 empty] | 1
| 1 | 2 | 2.3 | 2.15 | 1.86 | 1.9601 | 2.1 | 1.94 | 2.2 | 1.8 | 2.2453 | 1.9 | [19 empty] | 33
| 1 | 3 | 2.3 | 2.08 | 1.73 | 1.9776 | 2.1 | 1.95 | 2.1 | 1.8 | 2.112 | 1.9 | [19 empty] | 16
| 1 | 4 | 2.4 | 2.04 | 1.77 | 1.9476 | 2 | 2 | 2.2 | 1.8 | 2.0863 | 2 | [19 empty] | 2.0336
| 1 | 5 | 2.3 | 2.04 | 1.93 | 1.9374 | 2 | 2.01 | 2.1 | 1.9 | 2.2409 | 2 | [19 empty] | 70
| 1 | 6 | 2.4 | 2.29 | 1.81 | 1.9728 | 2.1 | 1.99 | 2.2 | 1.9 | 2.2589 | 2
| Mean |  | 2.3333 | 2.11 | 1.7983 | 1.9537 | 2.05 | 1.9867 | 2.15 | 1.8167 | 2.1901 | 1.9833
| Median |  | 2.3 | 2.07 | 1.79 | 1.9539 | 2.05 | 1.995 | 2.15 | 1.8 | 2.2189 | 2
| Std Dev. |  | 0.05164 | 0.09716 | 0.087731 | 0.020087 | 0.054772 | 0.035024 | 0.054772 | 0.075277 | 0.073847 | 0.075277
| Rel.Std.Dev. |  | 2.213% | 4.605% | 4.878% | 1.028% | 2.672% | 1.763% | 2.548% | 4.144% | 3.372% | 3.795%
| PDM3 |  | 14.74% | 3.757% | -11.57% | -3.93% | 0.8061% | -2.308% | 5.723% | -10.67% | 7.693% | -2.472%
| Z-Score (Absolute) |  | 1.9 | 0.55 | 1.33 | 0.39 | 0.19 | 0.19 | 0.8 | 1.22 | 1.04 | 0.21
| Analytical results for Zn in OREAS 234b (Certified Value 143 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 149 | 151 | 139.12 | 141.66 | 137.1196 | 136 | 157 | 147 | 142 | 132 | 135 | 157 | 158 | 157 | 151 | 144 | 146 | 140 | 143 | 101.9 | 137.4452 | 140 | 142.1899 | 152 | 184.8742 | 134 | 140 | 134 | 134 | 1
| 1 | 2 | 147 | 156 | 137.54 | 142.97 | 146.0351 | 136 | 148 | 146 | 148 | 132 | 135 | 152 | 153 | 148 | 152 | 143 | 149 | 143 | 137 | 61.42 | 131.6155 | 140 | 143.3067 | 150 | 181.0974 | 130 | 138 | 134 | 134 | 34
| 1 | 3 | 147 | 152 | 137.05 | 144.05 | 148.9631 | 136 | 148 | 147 | 144 | 133 | 142 | 155 | 154 | 152 | 152 | 143 | 148 | 142 | 136 | 62.9 | 129.7752 | 140 | 144.8743 | 153 | 183.8346 | 135 | 139 | 135 | 140 | 16
| 1 | 4 | 148 | 154 | 137 | 144.43 | 144.498 | 138 | 151 | 149 | 146 | 130 | 139 | 152 | 148 | 147 | 151 | 134 | 145 | 141 | 137 | 64.83 | 131.6058 | 140 | 150.0623 | 150 | 180.2499 | 134 | 142 | 134 | 142 | 143.0977
| 1 | 5 | 144 | 147 | 139.02 | 145.88 | 135.6738 | 135 | 155 | 149 | 145 | 137 | 138 | 148 | 149 | 157 | 154 | 142 | 145 | 140 | 144 | 62.14 | 130.6552 | 140 | 144.912 | 153 | 182.0349 | 137 | 140 | 133 | 136 | 71
| 1 | 6 | 149 | 146 | 138.02 | 146.76 | 143.6236 | 138 | 151 | 145 | 143 | 134 | 139 | 150 | 153 | 153 | 154 | 132 | 147 | 140 | 140 | 58.29 | 128.8431 | 140 | 143.5638 | 150 | 184.4245 | 135 | 146 | 134 | 151
| Mean |  | 147.3333 | 151 | 137.9583 | 144.2917 | 142.6522 | 136.5 | 151.6667 | 147.1667 | 144.6667 | 133 | 138 | 152.3333 | 152.5 | 152.3333 | 152.3333 | 139.6667 | 146.6667 | 141 | 139.5 | 68.58 | 131.6567 | 140 | 144.8182 | 151.3333 | 182.7526 | 134.1667 | 140.8333 | 134 | 139.5
| Median |  | 147.5 | 151.5 | 137.78 | 144.24 | 144.0608 | 136 | 151 | 147 | 144.5 | 132.5 | 138.5 | 152 | 153 | 152.5 | 152 | 142.5 | 146.5 | 140.5 | 138.5 | 62.52 | 131.1305 | 140 | 144.219 | 151 | 182.9348 | 134.5 | 140 | 134 | 138
| Std Dev. |  | 1.8619 | 3.8987 | 0.93796 | 1.8636 | 5.1948 | 1.2247 | 3.6697 | 1.6021 | 2.1602 | 2.3664 | 2.6833 | 3.266 | 3.6194 | 4.274 | 1.3663 | 5.2409 | 1.633 | 1.2649 | 3.3912 | 16.4631 | 3.0322 | 0 | 2.7669 | 1.5055 | 1.8966 | 2.3166 | 2.8577 | 0.63246 | 6.5038
| Rel.Std.Dev. |  | 1.264% | 2.582% | 0.6799% | 1.292% | 3.642% | 0.8972% | 2.42% | 1.089% | 1.493% | 1.779% | 1.944% | 2.144% | 2.373% | 2.806% | 0.8969% | 3.752% | 1.113% | 0.8971% | 2.431% | 24.01% | 2.303% | 0% | 1.911% | 0.9949% | 1.038% | 1.727% | 2.029% | 0.472% | 4.662%
| PDM3 |  | 2.96% | 5.522% | -3.591% | 0.8344% | -0.3113% | -4.611% | 5.988% | 2.844% | 1.096% | -7.056% | -3.562% | 6.454% | 6.571% | 6.454% | 6.454% | -2.398% | 2.494% | -1.466% | -2.514% | -52.07% | -7.995% | -2.165% | 1.202% | 5.755% | 27.71% | -6.241% | -1.582% | -6.358% | -2.514%
| Z-Score (Absolute) |  | 0.67 | 1.2 | 0.67 | 0.24 | 0 | 0.88 | 1.29 | 0.65 | 0.29 | 1.39 | 0.67 | 1.39 | 1.41 | 1.39 | 1.39 | 0.43 | 0.58 | 0.24 | 0.45 | 10.64 | 1.58 | 0.38 | 0.31 | 1.25 | 5.76 | 1.22 | 0.26 | 1.24 | 0.45
| Analytical results for Zr in OREAS 234b (Certified Value 49.6 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
1
| 1 | 1 | 51.8 | 52 | 73.75 | 56.8323 | 51.0573 | 45 | 59 | 52.2 | 44.5 | 53 | 58 | 50.5 | 50.4 | 49.1 | 47.9 | 43 | 39.9 | 44.8 | 47.5 | 51.0232 | 50 | 68.6554 | 52.9 | 33.5128 | 48.8 | 84.9 | 56 | 39.9 |  | 1
| 1 | 2 | 58.8 | 52.1 | 65.74 | 56.0951 | 50.6507 | 49 | 60 | 48.6 | 45.9 | 53 | 59.5 | 51.7 | 57.9 | 51.5 | 63.3 | 43 | 42.8 | 47 | 45.3 | 48.7968 | 50 | 68.6487 | 52.3 | 33.3959 | 46.9 | 79.8 | 57.4 | 39.4 |  | 35
| 1 | 3 | 60.1 | 46.9 | 73 | 55.5519 | 51.0086 | 57 | 57 | 48.9 | 43.7 | 53 | 57.4 | 50.7 | 50.5 | 52.7 | 46.3 | 43 | 40.5 | 45.8 | 44.6 | 47.5548 | 50 | 70.1749 | 55.4 | 33.161 | 47.7 | 82.3 | 56.1 | 41.2 |  | 16
| 1 | 4 | 57.1 | 49.1 | 65.89 | 54.7468 | 51.4139 | 52 | 59 | 48.3 | 44.8 | 54 | 56.5 | 48.5 | 50.5 | 56.7 | 46.3 | 41 | 42.9 | 43.2 | 46.1 | 48.2264 | 50 | 71.4956 | 55.2 | 32.3433 | 48 | 85 | 55 | 43.9 |  | 49.5548
| 1 | 5 | 58 | 48.7 | 65.59 | 56.0951 | 47.8854 | 56 | 58 | 48.7 | 45.4 | 54 | 57.8 | 48.4 | 51.2 | 58.3 | 46.2 | 44 | 43 | 43.8 | 48.3 | 48.1712 | 50 | 69.8697 | 55.5 | 32.6156 | 46.9 | 85.9 | 57.1 | 42.2 |  | 72
| 1 | 6 | 62.2 | 47.7 | 66.45 | 56.9778 | 70.3075 | 21 | 57 | 49.3 | 45 | 54 | 58.1 | 48.1 | 55.4 | 54.9 | 47.9 | 40 | 44.4 | 44.3 | 46.9 | 47.1132 | 50 | 69.2236 | 46.6 | 32.4974 | 48.4 | 84 | 58.1 | 47.8
| Mean |  | 58 | 49.4167 | 68.4033 | 56.0498 | 53.7206 | 46.6667 | 58.3333 | 49.3333 | 44.8833 | 53.5 | 57.8833 | 49.65 | 52.65 | 53.8667 | 49.65 | 42.3333 | 42.25 | 44.8167 | 46.45 | 48.4809 | 50 | 69.678 | 52.9833 | 32.921 | 47.7833 | 83.65 | 56.6167 | 42.4
| Median |  | 58.4 | 48.9 | 66.17 | 56.0951 | 51.033 | 50.5 | 58.5 | 48.8 | 44.9 | 53.5 | 57.9 | 49.5 | 50.85 | 53.8 | 47.1 | 43 | 42.85 | 44.55 | 46.5 | 48.1988 | 50 | 69.5467 | 54.05 | 32.8883 | 47.85 | 84.45 | 56.6 | 41.7
| Std Dev. |  | 3.5197 | 2.1803 | 3.8693 | 0.82744 | 8.2264 | 13.3367 | 1.2111 | 1.4431 | 0.75741 | 0.54772 | 0.98268 | 1.5043 | 3.2105 | 3.4186 | 6.7349 | 1.5055 | 1.7026 | 1.3891 | 1.3853 | 1.3749 | 0 | 1.0862 | 3.4149 | 0.49796 | 0.77825 | 2.2457 | 1.1232 | 3.1029
| Rel.Std.Dev. |  | 6.068% | 4.412% | 5.657% | 1.476% | 15.31% | 28.58% | 2.076% | 2.925% | 1.688% | 1.024% | 1.698% | 3.03% | 6.098% | 6.346% | 13.56% | 3.556% | 4.03% | 3.1% | 2.982% | 2.836% | 0% | 1.559% | 6.445% | 1.513% | 1.629% | 2.685% | 1.984% | 7.318%
| PDM3 |  | 17.04% | -0.2787% | 38.04% | 13.11% | 8.406% | -5.828% | 17.71% | -0.4469% | -9.427% | 7.961% | 16.81% | 0.1921% | 6.246% | 8.701% | 0.1921% | -14.57% | -14.74% | -9.561% | -6.265% | -2.167% | 0.8984% | 40.61% | 6.919% | -33.57% | -3.575% | 68.8% | 14.25% | -14.44%
| Z-Score (Absolute) |  | 1.23 | 0.06 | 2.79 | 0.93 | 0.58 | 0.47 | 1.28 | 0.07 | 0.74 | 0.55 | 1.21 | 0.03 | 0.42 | 0.61 | 0.03 | 1.12 | 1.14 | 0.75 | 0.51 | 0.2 | 0.03 | 2.98 | 0.47 | 2.54 | 0.31 | 5.08 | 1.02 | 1.11
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C930: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1042: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1132: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 234b (Certified Value 0.343 ppm) | [32 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.44 | 0.34 | 0.52 | 0.35588 | 0.379 | 0.37 | 0.32 | 0.34 | < 0.5 | 0.4 | 0.35 | 0.34 | 0.32 | < 10 | 0.34 | 0.2 | < 1 | 0.33 | 0.288 | 0.35 | 1.23 | 0.3 | 0.32189 | 0.36 | 0.42 | 0.32 | 0.35 | 0.34 | 0.36 |  |  | 1
| 1 | 2 | 0.45 | 0.33 | 0.51 | 0.36208 | 0.327 | 0.36 | 0.35 | 0.3 | < 0.5 | 0.4 | 0.34 | 0.34 | 0.35 | < 10 | 0.33 | 0.2 | < 1 | 0.31 | 0.316 | 0.33 | < 0.2 | 0.3 | 0.32681 | 0.35 | 0.336 | 0.35 | 0.37 | 0.33 | 0.38 |  |  | 19
| 1 | 3 | 0.47 | 0.34 | 0.52 | 0.36933 | 0.359 | 0.38 | 0.34 | 0.32 | < 0.5 | 0.3 | 0.37 | 0.34 | 0.35 | < 10 | 0.34 | 0.2 | < 1 | 0.32 | 0.309 | 0.33 | < 0.2 | 0.3 | 0.36579 | 0.34 | 0.45 | 0.31 | 0.36 | 0.34 | 0.36 |  |  | 16
| 1 | 4 | 0.47 | 0.33 | 0.5 | 0.36264 | 0.333 | 0.4 | 0.34 | 0.31 | < 0.5 | 0.3 | 0.36 | 0.34 | 0.34 | < 10 | 0.34 | 0.2 | < 1 | 0.33 | 0.316 | 0.35 | 0.23 | 0.3 | 0.3652 | 0.34 | 0.386 | 0.41 | 0.35 | 0.33 | 0.37 |  |  | 0.34318
| 1 | 5 | 0.48 | 0.35 | 0.5 | 0.35959 | 0.329 | 0.38 | 0.34 | 0.33 | < 0.5 | 0.3 | 0.34 | 0.34 | 0.32 | < 10 | 0.32 | < 0.2 | < 1 | 0.31 | 0.313 | 0.36 | 0.35 | 0.3 | 0.33502 | 0.35 | 0.364 | 0.33 | 0.36 | 0.35 | 0.35 |  |  | 74
| 1 | 6 | 0.46 | 0.35 | 0.52 | 0.36691 | 0.348 | 0.35 | 0.31 | 0.3 | < 0.5 | 0.3 | 0.34 | 0.34 | 0.35 | < 10 | 0.33 | 0.2 | < 1 | 0.33 | 0.331 | 0.36 | 0.24 | 0.3 | 0.33067 | 0.35 | 0.36 | 0.3 | 0.37 | 0.34 | 0.34
| Mean |  | 0.46167 | 0.34 | 0.51167 | 0.36274 | 0.34583 | 0.37333 | 0.33333 | 0.31667 |  | 0.33333 | 0.35 | 0.34 | 0.33833 |  | 0.33333 | 0.2 |  | 0.32167 | 0.31217 | 0.34667 | 0.5125 | 0.3 | 0.3409 | 0.34833 | 0.386 | 0.33667 | 0.36 | 0.33833 | 0.36
| Median |  | 0.465 | 0.34 | 0.515 | 0.36236 | 0.3405 | 0.375 | 0.34 | 0.315 |  | 0.3 | 0.345 | 0.34 | 0.345 |  | 0.335 | 0.2 |  | 0.325 | 0.3145 | 0.35 | 0.295 | 0.3 | 0.33285 | 0.35 | 0.375 | 0.325 | 0.36 | 0.34 | 0.36
| Std Dev. |  | 0.01472 | 0.0089443 | 0.0098319 | 0.0048619 | 0.0204 | 0.017512 | 0.015055 | 0.01633 |  | 0.05164 | 0.012649 | 0 | 0.01472 |  | 0.008165 | 0 |  | 0.0098319 | 0.013992 | 0.013663 | 0.48141 | 0 | 0.01954 | 0.0075277 | 0.042218 | 0.039833 | 0.0089443 | 0.0075277 | 0.014142
| Rel.Std.Dev. |  | 3.188% | 2.631% | 1.922% | 1.34% | 5.899% | 4.691% | 4.517% | 5.157% |  | 15.49% | 3.614% | 0% | 4.351% |  | 2.449% | 0% |  | 3.057% | 4.482% | 3.941% | 93.93% | 0% | 5.732% | 2.161% | 10.94% | 11.83% | 2.485% | 2.225% | 3.928%
| PDM3 |  | 34.53% | -0.9274% | 49.09% | 5.698% | 0.7724% | 8.786% | -2.87% | -7.726% |  | -2.87% | 1.987% | -0.9274% | -1.413% |  | -2.87% | -41.72% |  | -6.27% | -9.038% | 1.015% | 49.34% | -12.58% | -0.6661% | 1.501% | 12.48% | -1.899% | 4.9% | -1.413% | 4.9%
| Z-Score (Absolute) |  | 5.51 | 0.28 | 7.89 | 0.8 | 0 | 1.31 | 0.59 | 1.39 | 4.56 | NA | 0.2 | 0.28 | 0.36 | 221.53 | 0.59 | 7.73 | 7.34 | 1.15 | 1.6 | 0.04 | 1.39 | NA | 0.23 | 0.12 | 1.91 | 0.44 | 0.67 | 0.36 | 0.67
| Analytical results for Al in OREAS 234b (Certified Value 3.12 wt.%) | [32 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.14 | 2.894 | 2.7105 | 3.5537 | 3.1611 | 3.06 | 3.525 | 3.6 | 2.9494 | 2.97 | 3.23 | 3.24 | 3.18 | 3.16 | 3.29 | 3.13 | 3.5089 | 3.307 | 2.7 | 3.32 | 3.145 | 2.8659 | 2.745 | 3.03 | 2.8 | 3.38 | 3.02 | 2.77 |  |  |  | 1
| 1 | 2 | 3.12 | 2.902 | 2.7069 | 3.5815 | 3.2812 | 2.84 | 3.456 | 3.52 | 2.8922 | 2.94 | 3.3 | 3.25 | 3.04 | 3.21 | 3.16 | 3.22 | 3.5086 | 3.549 | 3.11 | 3.34 | 2.9909 | 2.7918 | 2.687 | 3.02 | 2.92 | 3.34 | 3.01 | 2.76 |  |  |  | #N/A
| 1 | 3 | 3.14 | 2.883 | 2.813 | 3.5531 | 3.2654 | 2.87 | 3.398 | 3.5 | 2.9475 | 2.88 | 3.27 | 3.2 | 3.08 | 3.22 | 3.21 | 3.22 | 3.5073 | 3.382 | 3.1 | 3.38 | 3.3952 | 2.6494 | 2.77 | 3.05 | 2.86 | 3.35 | 3.04 | 2.79 |  |  |  | 16
| 1 | 4 | 3.21 | 2.879 | 2.7197 | 3.5525 | 3.1845 | 2.82 | 3.424 | 3.46 | 2.9371 | 2.94 | 3.26 | 3.17 | 3.11 | 3.21 | 3.21 | 3.06 | 3.3969 | 3.621 | 2.81 | 3.42 | 2.9468 | 2.7148 | 2.756 | 3.05 | 2.9 | 3.36 | 3.02 | 2.79 |  |  |  | 3.1165
| 1 | 5 | 3.2 | 2.937 | 2.7853 | 3.5648 | 3.0669 | 2.81 | 3.234 | 3.56 | 3.0418 | 2.92 | 3.27 | 3.2 | 3.08 | 3.14 | 3.22 | 3.08 | 3.6318 | 3.325 | 2.82 | 3.37 | 3.0457 | 2.8339 | 2.6765 | 3 | 2.8 | 3.34 | 3.12 | 2.82 |  |  |  | 75
| 1 | 6 | 3.11 | 2.907 | 2.7026 | 3.5781 | 3.0988 | 2.93 | 3.445 | 3.5 | 2.8994 | 2.93 | 3.27 | 3.09 | 3.06 | 3.15 | 3.18 | 3.3 | 3.4858 | 3.5 | 3.03 | 3.37 | 3.3242 | 2.8346 | 2.642 | 2.99 | 2.79 | 3.35 | 3.05 | 2.77
| Mean |  | 3.1533 | 2.9003 | 2.7397 | 3.5639 | 3.1763 | 2.8883 | 3.4137 | 3.5233 | 2.9446 | 2.93 | 3.2667 | 3.1917 | 3.0917 | 3.1817 | 3.2117 | 3.1683 | 3.5066 | 3.4473 | 2.9283 | 3.3667 | 3.1413 | 2.7817 | 2.7127 | 3.0233 | 2.845 | 3.3533 | 3.0433 | 2.7833
| Median |  | 3.14 | 2.898 | 2.7151 | 3.5593 | 3.1728 | 2.855 | 3.4345 | 3.51 | 2.9423 | 2.935 | 3.27 | 3.2 | 3.08 | 3.185 | 3.21 | 3.175 | 3.5079 | 3.441 | 2.925 | 3.37 | 3.0953 | 2.8128 | 2.716 | 3.025 | 2.83 | 3.35 | 3.03 | 2.78
| Std Dev. |  | 0.041793 | 0.020916 | 0.047228 | 0.013136 | 0.086305 | 0.094534 | 0.097777 | 0.049666 | 0.053524 | 0.029665 | 0.022509 | 0.057764 | 0.04916 | 0.035449 | 0.04446 | 0.09347 | 0.075039 | 0.12822 | 0.1736 | 0.034448 | 0.18303 | 0.083285 | 0.051324 | 0.025033 | 0.05648 | 0.015055 | 0.040332 | 0.021602
| Rel.Std.Dev. |  | 1.325% | 0.7211% | 1.724% | 0.3686% | 2.717% | 3.273% | 2.864% | 1.41% | 1.818% | 1.012% | 0.6891% | 1.81% | 1.59% | 1.114% | 1.384% | 2.95% | 2.14% | 3.719% | 5.928% | 1.023% | 5.826% | 2.994% | 1.892% | 0.828% | 1.985% | 0.449% | 1.325% | 0.7761%
| PDM3 |  | 1.182% | -6.936% | -12.09% | 14.36% | 1.919% | -7.321% | 9.535% | 13.05% | -5.517% | -5.984% | 4.818% | 2.412% | -0.797% | 2.091% | 3.054% | 1.663% | 12.52% | 10.62% | -6.038% | 8.027% | 0.7957% | -10.74% | -12.96% | -2.99% | -8.712% | 7.599% | -2.348% | -10.69%
| Z-Score (Absolute) |  | 0.02 | 0.76 | 1.26 | 1.29 | 0.09 | 0.8 | 0.82 | 1.16 | 0.63 | 0.67 | 0.37 | 0.14 | 0.17 | 0.11 | 0.2 | 0.06 | 1.11 | 0.93 | 0.68 | 0.68 | 0.02 | 1.13 | 1.34 | 0.38 | 0.93 | 0.64 | 0.32 | 1.13
| Analytical results for As in OREAS 234b (Certified Value 65 ppm) | [32 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 1
| 1 | 1 | 68 | 64.77 | 56.31 | 62.2296 | 64.3737 | 67.4 | 68.9 | 62 | 64 | 61 | 62.2 | 67.4 | 67.1 | 63.2 | < 100 | 68.7 | 62 | 52.2293 | 63 | 55.2 | 64 | 48.17 | 68.2272 | 65 | 65.7961 | 68.6 | 59.405 | 66.2 | 66 | 67.4 | 62.8 | 1
| 1 | 2 | 67 | 64.33 | 54.15 | 63.6846 | 63.8714 | 62.6 | 66.7 | 67 | 63 | 60 | 64.3 | 66.6 | 70.2 | 66.6 | < 100 | 67 | 60 | 50.1209 | 61 | 60.9 | 63 | 48.59 | 67.4912 | 60 | 67.9562 | 68 | 59.276 | 67.9 | 67 | 65.9 | 65.2 | 20
| 1 | 3 | 68 | 63.88 | 59.59 | 64.1688 | 64.8204 | 65 | 68.3 | 66 | 63 | 58 | 62.6 | 67.4 | 69.4 | 69.2 | < 100 | 68.3 | 62 | 50.7996 | 63 | 65.4 | 64 | 48.92 | 63.4708 | 65 | 71.0539 | 67.8 | 63.598 | 65.1 | 65 | 67 | 63.5 | 16
| 1 | 4 | 69 | 64.93 | 55.83 | 63.57 | 63.7139 | 61.9 | 68.1 | 66 | 64 | 59 | 63.5 | 66.7 | 68 | 66.3 | < 100 | 67.5 | 59 | 53.7988 | 62 | 60.1 | 63 | 49.14 | 64.1792 | 65 | 71.8135 | 68.5 | 62.519 | 68.4 | 64 | 64.7 | 65.5 | 64.9188
| 1 | 5 | 68 | 64.37 | 54.9 | 63.589 | 65.2208 | 62.8 | 71.5 | 66 | 65 | 59 | 61.8 | 66 | 67.8 | 65.5 | < 100 | 67.4 | 58 | 56.3741 | 61 | 59.2 | 64 | 48.99 | 66.4056 | 65 | 69.6593 | 67.5 | 61.546 | 65.2 | 66 | 67.5 | 64.6 | 76
| 1 | 6 | 67 | 64.08 | 55.09 | 62.5347 | 65.7384 | 63 | 68.3 | 67 | 63 | 59 | 61.3 | 66.7 | 68.5 | 66 | < 100 | 66.2 | 63 | 54.9175 | 61 | 61 | 64 | 48.21 | 65.918 | 65 | 67.8149 | 67.6 | 68.593 | 66.5 | 67 | 65 | 62.4
| Mean |  | 67.8333 | 64.3933 | 55.9783 | 63.2961 | 64.6231 | 63.7833 | 68.6333 | 65.6667 | 63.6667 | 59.3333 | 62.6167 | 66.8 | 68.5 | 66.1333 |  | 67.5167 | 60.6667 | 53.04 | 61.8333 | 60.3 | 63.6667 | 48.67 | 65.9487 | 64.1667 | 69.0156 | 68 | 62.4895 | 66.55 | 65.8333 | 66.25 | 64
| Median |  | 68 | 64.35 | 55.46 | 63.5795 | 64.5971 | 62.9 | 68.3 | 66 | 63.5 | 59 | 62.4 | 66.7 | 68.25 | 66.15 |  | 67.45 | 61 | 53.014 | 61.5 | 60.5 | 64 | 48.755 | 66.1618 | 65 | 68.8077 | 67.9 | 62.0325 | 66.35 | 66 | 66.45 | 64.05
| Std Dev. |  | 0.75277 | 0.39903 | 1.922 | 0.74704 | 0.78666 | 2.0537 | 1.5832 | 1.8619 | 0.8165 | 1.0328 | 1.1125 | 0.53292 | 1.1314 | 1.9325 |  | 0.89759 | 1.9664 | 2.4255 | 0.98319 | 3.29 | 0.5164 | 0.41323 | 1.8467 | 2.0412 | 2.2506 | 0.46043 | 3.4413 | 1.3635 | 1.169 | 1.2276 | 1.2884
| Rel.Std.Dev. |  | 1.11% | 0.6197% | 3.434% | 1.18% | 1.217% | 3.22% | 2.307% | 2.835% | 1.282% | 1.741% | 1.777% | 0.7978% | 1.652% | 2.922% |  | 1.329% | 3.241% | 4.573% | 1.59% | 5.456% | 0.8111% | 0.849% | 2.8% | 3.181% | 3.261% | 0.6771% | 5.507% | 2.049% | 1.776% | 1.853% | 2.013%
| PDM3 |  | 4.49% | -0.8094% | -13.77% | -2.5% | -0.4554% | -1.749% | 5.722% | 1.152% | -1.929% | -8.604% | -3.546% | 2.898% | 5.517% | 1.871% |  | 4.002% | -6.55% | -18.3% | -4.753% | -7.115% | -1.929% | -25.03% | 1.586% | -1.159% | 6.311% | 4.746% | -3.742% | 2.513% | 1.409% | 2.051% | -1.415%
| Z-Score (Absolute) |  | 1.06 | 0.07 | 2.37 | 0.25 | 0.13 | 0.11 | 1.29 | 0.43 | 0.14 | 1.4 | 0.45 | 0.76 | 1.25 | 0.57 | 4.09 | 0.97 | 1.01 | 3.22 | 0.67 | 1.12 | 0.14 | 4.48 | 0.51 | 0 | 1.4 | 1.11 | 0.48 | 0.69 | 0.48 | 0.6 | 0.05
| Analytical results for B in OREAS 234b (Certified Value 26.9 ppm) | [32 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.11 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 22 | 22 | 35.06 | 27.6148 | 42 | 32 | 20 | 20 | 20 | 20 | 25 | 17 | 58.27 | 20 | 20 | 15 | < 10 | 32 | 17.8 | [12 empty] | 1
| 1 | 2 | 27 | 25 | 39.04 | 27.5478 | 38 | 32 | 20 | 20 | 20 | 20 | 25 | 23 | 61.55 | 20 | 10 | 16 | < 10 | 34 | 18 | [12 empty] | 3
| 1 | 3 | 25 | 23 | 37.26 | 26.9869 | 38 | 30 | 20 | 20 | 30 | 20 | 27 | 22 | 60.25 | 20 | 10 | 14 | < 10 | 33 | 17.6 | [12 empty] | 16
| 1 | 4 | 23 | 22 | 37.1 | 27.5996 | 39 | 32 | 20 | 20 | 20 | 20 | 26 | 18 | 60.14 | 20 | 20 | 16 | < 10 | 32 | 18.4 | [12 empty] | 26.851
| 1 | 5 | 28 | 25 | 36.85 | 26.5623 | 39 | 33 | 20 | 20 | 20 | 20 | 28 | 21 | 59.57 | 20 | 10 | 15 | < 10 | 35 | 21 | [12 empty] | 77
| 1 | 6 | 25 | 24 | 36.4 | 26.2273 | 38 | 31 | 20 | 20 | 20 | 20 | 28 | 21 | 58.03 | 20 | 10 | 14 | < 10 | 32 | 17.8
| Mean |  | 25 | 23.5 | 36.9517 | 27.0898 | 39 | 31.6667 | 20 | 20 | 21.6667 | 20 | 26.5 | 20.3333 | 59.635 | 20 | 13.3333 | 15 |  | 33 | 18.4333
| Median |  | 25 | 23.5 | 36.975 | 27.2673 | 38.5 | 32 | 20 | 20 | 20 | 20 | 26.5 | 21 | 59.855 | 20 | 10 | 15 |  | 32.5 | 17.9
| Std Dev. |  | 2.2804 | 1.3784 | 1.2939 | 0.59634 | 1.5492 | 1.0328 | 0 | 0 | 4.0825 | 0 | 1.3784 | 2.3381 | 1.3224 | 0 | 5.164 | 0.89443 |  | 1.2649 | 1.2863
| Rel.Std.Dev. |  | 9.121% | 5.866% | 3.502% | 2.201% | 3.972% | 3.261% | 0% | 0% | 18.84% | 0% | 5.202% | 11.5% | 2.218% | 0% | 38.73% | 5.963% |  | 3.833% | 6.978%
| PDM3 |  | -6.894% | -12.48% | 37.62% | 0.8891% | 45.25% | 17.93% | -25.51% | -25.51% | -19.31% | -25.51% | -1.307% | -24.27% | 122.1% | -25.51% | -50.34% | -44.14% |  | 22.9% | -31.35%
| Z-Score (Absolute) |  | 0.16 | 0.31 | 1.08 | 0.06 | 1.3 | 0.54 | NA | NA | NA | NA | 0 | 0.64 | 3.44 | NA | NA | 1.19 | 2.23 | 0.67 | 0.84
Results from laboratories 1.14, 1.15, 1.16, 1.18, 1.27 and 1.29 were removed due to their 10 ppm reading resolution.
| Analytical results for Ba in OREAS 234b (Certified Value 41.2 ppm) | [32 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 41 | 41.66 | 41.18 | 47.47 | 42.3622 | 40.2 | 51 | 47 | 41 | 42 | 41 | 40 | 40 | 40 | 56.553 | 40 | 29 | 45.7017 | 42 | 30.8 | 40 | 35.14 | 40.707 | 34 | 40 | 39 | 45 | 38 | 37 |  |  | 1
| 1 | 2 | 40 | 41.27 | 41.32 | 47.85 | 42.6867 | 39.3 | 50 | 49 | 40 | 42 | 42 | 40 | 40 | 40 | 57.118 | 40 | 29 | 44.7466 | 43 | 34.6 | 41 | 36.88 | 39.438 | 33 | 40 | 40 | 44 | 38 | 37 |  |  | 21
| 1 | 3 | 41 | 40.93 | 42.75 | 48.93 | 42.5971 | 40.7 | 49 | 47 | 41 | 41 | 41 | 40 | 40 | 40 | 54.757 | 40 | 29 | 45.305 | 43 | 33.9 | 41 | 37.28 | 37.782 | 34 | 40 | 40 | 45 | 38 | 37 |  |  | 16
| 1 | 4 | 41 | 42.42 | 41.65 | 45.72 | 41.9701 | 38.4 | 50 | 47 | 41 | 41 | 42 | 40 | 40 | 40 | 56.738 | 40 | 27 | 44.2998 | 47 | 31.1 | 41 | 38.09 | 39.06 | 33 | 40 | 40 | 44 | 38 | 38 |  |  | 41.1843
| 1 | 5 | 51 | 41.91 | 40.5 | 48.24 | 43.2127 | 39.9 | 48 | 47 | 42 | 42 | 41 | 40 | 40 | 40 | 57.692 | 40 | 29 | 45.6509 | 43 | 31.3 | 40 | 36.77 | 40.428 | 34 | 40 | 39 | 45 | 39 | 37 |  |  | 78
| 1 | 6 | 40 | 41.3 | 41.16 | 47 | 43.1517 | 41.6 | 49 | 48 | 40 | 41 | 41 | 40 | 40 | 40 | 59.59 | 40 | 30 | 45.7511 | 48 | 33.7 | 41 | 36.36 | 40.185 | 31 | 40 | 39 | 46 | 38 | 36
| Mean |  | 42.3333 | 41.5817 | 41.4267 | 47.535 | 42.6634 | 40.0167 | 49.5 | 47.5 | 40.8333 | 41.5 | 41.3333 | 40 | 40 | 40 | 57.0747 | 40 | 28.8333 | 45.2425 | 44.3333 | 32.5667 | 40.6667 | 36.7533 | 39.6 | 33.1667 | 40 | 39.5 | 44.8333 | 38.1667 | 37
| Median |  | 41 | 41.48 | 41.25 | 47.66 | 42.6419 | 40.05 | 49.5 | 47 | 41 | 41.5 | 41 | 40 | 40 | 40 | 56.928 | 40 | 29 | 45.4779 | 43 | 32.5 | 41 | 36.825 | 39.8115 | 33.5 | 40 | 39.5 | 45 | 38 | 37
| Std Dev. |  | 4.274 | 0.53252 | 0.74875 | 1.107 | 0.47255 | 1.1089 | 1.0488 | 0.83666 | 0.75277 | 0.54772 | 0.5164 | 0 | 0 | 0 | 1.579 | 0 | 0.98319 | 0.59578 | 2.5033 | 1.6777 | 0.5164 | 0.98366 | 1.0841 | 1.169 | 0 | 0.54772 | 0.75277 | 0.40825 | 0.63246
| Rel.Std.Dev. |  | 10.1% | 1.281% | 1.807% | 2.329% | 1.108% | 2.771% | 2.119% | 1.761% | 1.844% | 1.32% | 1.249% | 0% | 0% | 0% | 2.767% | 0% | 3.41% | 1.317% | 5.647% | 5.152% | 1.27% | 2.676% | 2.738% | 3.525% | 0% | 1.387% | 1.679% | 1.07% | 1.709%
| PDM3 |  | 2.79% | 0.9649% | 0.5885% | 15.42% | 3.591% | -2.835% | 20.19% | 15.34% | -0.8522% | 0.7666% | 0.3619% | -2.876% | -2.876% | -2.876% | 38.58% | -2.876% | -29.99% | 9.854% | 7.646% | -20.92% | -1.257% | -10.76% | -3.847% | -19.47% | -2.876% | -4.09% | 8.86% | -7.327% | -10.16%
| Z-Score (Absolute) |  | 0.21 | 0.04 | 0.01 | 1.35 | 0.28 | 0.3 | 1.78 | 1.34 | 0.12 | 0.03 | 0.01 | NA | NA | NA | 3.43 | NA | 2.74 | 0.84 | 0.65 | 1.93 | 0.16 | 1.01 | 0.39 | 1.8 | NA | 0.41 | 0.76 | 0.7 | 0.96
Results from laboratories 1.14, 1.15, 1.16, 1.18 and 1.29 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 234b (Certified Value 0.21 ppm) | [32 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.3 | 0.22 | < 1 | 0.14535 | 0.23699 | 0.3 | 0.23 | < 0.5 | 0.25 | < 5 | < 1 | 0.22 | 0.21 | 0.2 | < 10 | 0.22 | < 0.5 | 1.035 | < 0.5 | < 0.5 | 0.2 | 0.19 | 0.09 | 0.3 | 0.24 | 0.14 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.26 | < 1 | 0.14839 | 0.2286 | 0.2 | 0.26 | < 0.5 | 0.24 | < 5 | < 1 | 0.22 | 0.21 | 0.21 | < 10 | 0.21 | < 0.5 | 1.0023 | < 0.5 | < 0.5 | 0.2 | 0.19 | 0.11 | 0.3 | 0.26 | 0.18 | [5 empty] | 22
| 1 | 3 | 0.2 | 0.24 | < 1 | 0.14622 | 0.23504 | 0.2 | 0.23 | < 0.5 | 0.22 | < 5 | < 1 | 0.22 | 0.2 | 0.21 | < 10 | 0.21 | < 0.5 | 1.0031 | < 0.5 | < 0.5 | 0.2 | 0.2 | 0.11 | 0.3 | 0.25 | 0.16 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.22 | < 1 | 0.14627 | 0.24107 | 0.2 | 0.31 | < 0.5 | 0.22 | < 5 | < 1 | 0.22 | 0.21 | 0.21 | < 10 | 0.22 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.2 | < 0.05 | 0.3 | 0.24 | 0.17 | [5 empty] | 0.21139
| 1 | 5 | 0.2 | 0.24 | < 1 | 0.14334 | 0.23704 | 0.2 | 0.29 | < 0.5 | 0.23 | < 5 | < 1 | 0.22 | 0.21 | 0.2 | < 10 | 0.22 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.19 | 0.08 | 0.3 | 0.24 | 0.14 | [5 empty] | 79
| 1 | 6 | 0.2 | 0.23 | < 1 | 0.14366 | 0.23881 | 0.2 | 0.17 | < 0.5 | 0.22 | < 5 | < 1 | 0.2 | 0.21 | 0.2 | < 10 | 0.22 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.19 | 0.05 | 0.3 | 0.23 | 0.15
| Mean |  | 0.21667 | 0.235 |  | 0.14554 | 0.23626 | 0.21667 | 0.24833 |  | 0.23 |  |  | 0.21667 | 0.20833 | 0.205 |  | 0.21667 |  | 1.0135 |  |  | 0.2 | 0.19333 | 0.088 | 0.3 | 0.24333 | 0.15667
| Median |  | 0.2 | 0.235 |  | 0.14579 | 0.23701 | 0.2 | 0.245 |  | 0.225 |  |  | 0.22 | 0.21 | 0.205 |  | 0.22 |  | 1.0031 |  |  | 0.2 | 0.19 | 0.09 | 0.3 | 0.24 | 0.155
| Std Dev. |  | 0.040825 | 0.015166 |  | 0.0018735 | 0.0042644 | 0.040825 | 0.049967 |  | 0.012649 |  |  | 0.008165 | 0.0040825 | 0.0054772 |  | 0.005164 |  | 0.018666 |  |  | 3.0405e-17 | 0.005164 | 0.0249 | 0 | 0.010328 | 0.01633
| Rel.Std.Dev. |  | 18.84% | 6.454% |  | 1.287% | 1.805% | 18.84% | 20.12% |  | 5.5% |  |  | 3.768% | 1.96% | 2.672% |  | 2.383% |  | 1.842% |  |  | 1.52e-14% | 2.671% | 28.3% | 0% | 4.244% | 10.42%
| PDM3 |  | 2.497% | 11.17% |  | -31.15% | 11.76% | 2.497% | 17.48% |  | 8.804% |  |  | 2.497% | -1.445% | -3.022% |  | 2.497% |  | 379.4% |  |  | -5.388% | -8.541% | -58.37% | 41.92% | 15.11% | -25.89%
| Z-Score (Absolute) |  | NA | 0.12 | 6.43 | 2.33 | 0.09 | NA | 0.21 | 0.25 | 0.24 | 55.84 | 6.43 | 0.57 | 0.78 | 0.86 | 117.59 | 0.57 | 0.25 | 12.77 | 0.25 | 0.25 | NA | 1.15 | 4.01 | NA | 0.09 | 2.05
Results from laboratories 1.01, 1.07, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 234b (Certified Value 0.042 ppm) | [32 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.02 | 0.05 | < 1 | 0.021651 | 0.08 | 0.07 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.04 | 0.05 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.05 | < 0.1 | 0.04401 | 0.04 | 0.035 | 0.06 | 0.05 | 0.05 | 0.03 |  |  |  |  | 1
| 1 | 2 | 0.04 | 0.04 | < 1 | 0.026842 | 0.04 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.05 | 0.04 | < 10 | 0.04 | < 5 | < 0.05 | 0.04 | 0.05 | < 0.1 | 0.05185 | 0.04 | 0.032 | 0.05 | 0.05 | 0.05 | 0.03 |  |  |  |  | 23
| 1 | 3 | 0.02 | 0.05 | < 1 | 0.026659 | 0.04 | 0.07 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.04 | 0.05 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.04 | < 0.1 | 0.05286 | 0.04 | 0.027 | 0.05 | 0.05 | 0.05 | 0.03 |  |  |  |  | 16
| 1 | 4 | 0.03 | 0.05 | < 1 | 0.020944 | 0.04 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.04 | 0.05 | < 10 | 0.05 | < 5 | < 0.05 | 0.03 | 0.04 | < 0.1 | 0.04675 | 0.04 | 0.035 | 0.06 | 0.05 | 0.06 | 0.03 |  |  |  |  | 0.042285
| 1 | 5 | 0.02 | 0.04 | < 1 | 0.029468 | 0.04 | 0.06 | < 0.05 | 0.04 | < 2 | < 0.1 | 0.04 | 0.05 | 0.05 | < 10 | 0.04 | < 5 | < 0.05 | 0.03 | 0.05 | < 0.1 | 0.04566 | 0.04 | 0.029 | 0.05 | 0.05 | 0.05 | 0.03 |  |  |  |  | 80
| 1 | 6 | 0.03 | 0.04 | < 1 | 0.028362 | 0.05 | 0.05 | < 0.05 | 0.05 | < 2 | < 0.1 | 0.05 | 0.04 | 0.05 | < 10 | 0.05 | < 5 | < 0.05 | 0.03 | 0.05 | < 0.1 | 0.05638 | 0.04 | 0.029 | 0.05 | 0.05 | 0.05 | 0.03
| Mean |  | 0.026667 | 0.045 |  | 0.025654 | 0.048333 | 0.061667 |  | 0.041667 |  |  | 0.041667 | 0.043333 | 0.048333 |  | 0.043333 |  |  | 0.031667 | 0.046667 |  | 0.049585 | 0.04 | 0.031167 | 0.053333 | 0.05 | 0.051667 | 0.03
| Median |  | 0.025 | 0.045 |  | 0.02675 | 0.04 | 0.06 |  | 0.04 |  |  | 0.04 | 0.04 | 0.05 |  | 0.04 |  |  | 0.03 | 0.05 |  | 0.0493 | 0.04 | 0.0305 | 0.05 | 0.05 | 0.05 | 0.03
| Std Dev. |  | 0.008165 | 0.0054772 |  | 0.0035358 | 0.016021 | 0.0075277 |  | 0.0040825 |  |  | 0.0040825 | 0.005164 | 0.0040825 |  | 0.005164 |  |  | 0.0040825 | 0.005164 |  | 0.0048281 | 0 | 0.0033714 | 0.005164 | 7.6012e-18 | 0.0040825 | 0
| Rel.Std.Dev. |  | 30.62% | 12.17% |  | 13.78% | 33.15% | 12.21% |  | 9.798% |  |  | 9.798% | 11.92% | 8.447% |  | 11.92% |  |  | 12.89% | 11.07% |  | 9.737% | 0% | 10.82% | 9.682% | 1.52e-14% | 7.902% | 0%
| PDM3 |  | -36.94% | 6.422% |  | -39.33% | 14.31% | 45.84% |  | -1.461% |  |  | -1.461% | 2.48% | 14.31% |  | 2.48% |  |  | -25.11% | 10.36% |  | 17.27% | -5.403% | -26.29% | 26.13% | 18.25% | 22.19% | -29.05%
| Z-Score (Absolute) |  | 2.02 | 0.17 | 45.85 | 2.13 | 0.17 | 1.52 | 2.19 | 0.51 | 96.43 | 0.34 | 0.51 | 0.34 | 0.17 | 501.01 | 0.34 | 248.15 | 2.19 | 1.52 | 0 | 0.34 | 0.3 | 0.67 | 1.57 | 0.67 | 0.34 | 0.51 | 1.69
| Analytical results for Ca in OREAS 234b (Certified Value 2.2 wt.%) | [32 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 2.18 | 2.191 | 2.3492 | 2.7262 | 2.382 | 2.93 | 2.659 | 2.98 | 2.51 | 2.37 | 1.96 | 2.28 | 2.27 | 2.13 | 1.6507 | 2.28 | 2.15 | 2.8168 | 2.26 | 1.85 | 2.18 | 1.5991 | 2.4293 | 1.5675 | 2.18 | 1.74 | 2.06 | 2.12 | 1.74 |  |  | 1
| 1 | 2 | 2.16 | 2.198 | 2.4065 | 2.7735 | 2.3294 | 2.77 | 2.616 | 2.89 | 2.44 | 2.38 | 2 | 2.29 | 2.12 | 2.19 | 1.6227 | 2.24 | 2.2 | 2.7316 | 2.59 | 2.12 | 2.09 | 1.5728 | 2.3258 | 1.533 | 2.16 | 1.86 | 2.03 | 2.15 | 1.67 |  |  | #N/A
| 1 | 3 | 2.21 | 2.181 | 2.4171 | 2.7755 | 2.389 | 2.84 | 2.53 | 2.9 | 2.47 | 2.35 | 2 | 2.28 | 2.16 | 2.21 | 1.6218 | 2.26 | 2.16 | 2.6658 | 2.38 | 2.1 | 2.12 | 1.7245 | 2.3063 | 1.5715 | 2.15 | 1.86 | 2.05 | 2.16 | 1.7 |  |  | 16
| 1 | 4 | 2.27 | 2.19 | 2.4264 | 2.7621 | 2.3385 | 2.76 | 2.566 | 2.88 | 2.49 | 2.38 | 2 | 2.23 | 2.15 | 2.18 | 1.6518 | 2.27 | 2.01 | 2.707 | 2.7 | 1.9 | 2.16 | 1.597 | 2.2928 | 1.5755 | 2.17 | 1.86 | 2.09 | 2.15 | 1.71 |  |  | 2.201
| 1 | 5 | 2.29 | 2.209 | 2.3488 | 2.741 | 2.3973 | 2.82 | 2.394 | 2.96 | 2.56 | 2.34 | 2 | 2.27 | 2.18 | 2.12 | 1.665 | 2.32 | 2.11 | 2.8544 | 2.36 | 1.92 | 2.21 | 1.6319 | 2.3947 | 1.4965 | 2.12 | 1.78 | 2.05 | 2.23 | 1.75 |  |  | 81
| 1 | 6 | 2.21 | 2.171 | 2.3257 | 2.7459 | 2.3956 | 2.75 | 2.63 | 2.9 | 2.45 | 2.36 | 2 | 2.14 | 2.15 | 2.16 | 1.7067 | 2.22 | 2.25 | 2.8289 | 2.5 | 2.1 | 2.25 | 1.7243 | 2.4466 | 1.502 | 2.13 | 1.72 | 2.05 | 2.18 | 1.73
| Mean |  | 2.22 | 2.19 | 2.3789 | 2.754 | 2.372 | 2.8117 | 2.5658 | 2.9183 | 2.4867 | 2.3633 | 1.9933 | 2.2483 | 2.1717 | 2.165 | 1.6531 | 2.265 | 2.1467 | 2.7674 | 2.465 | 1.9983 | 2.1683 | 1.6416 | 2.3659 | 1.541 | 2.1517 | 1.8033 | 2.055 | 2.165 | 1.7167
| Median |  | 2.21 | 2.1905 | 2.3779 | 2.754 | 2.3855 | 2.795 | 2.591 | 2.9 | 2.48 | 2.365 | 2 | 2.275 | 2.155 | 2.17 | 1.6513 | 2.265 | 2.155 | 2.7742 | 2.44 | 2.01 | 2.17 | 1.6155 | 2.3602 | 1.5503 | 2.155 | 1.82 | 2.05 | 2.155 | 1.72
| Std Dev. |  | 0.050596 | 0.013176 | 0.042644 | 0.019573 | 0.030057 | 0.067946 | 0.096047 | 0.041191 | 0.044121 | 0.01633 | 0.01633 | 0.057067 | 0.051929 | 0.035071 | 0.031386 | 0.034496 | 0.082138 | 0.076202 | 0.1627 | 0.12106 | 0.058452 | 0.066824 | 0.066129 | 0.035744 | 0.023166 | 0.065013 | 0.019748 | 0.037283 | 0.029439
| Rel.Std.Dev. |  | 2.279% | 0.6016% | 1.793% | 0.7107% | 1.267% | 2.417% | 3.743% | 1.411% | 1.774% | 0.691% | 0.8192% | 2.538% | 2.391% | 1.62% | 1.899% | 1.523% | 3.826% | 2.754% | 6.6% | 6.058% | 2.696% | 4.071% | 2.795% | 2.32% | 1.077% | 3.605% | 0.961% | 1.722% | 1.715%
| PDM3 |  | 0.8653% | -0.4978% | 8.087% | 25.13% | 7.77% | 27.75% | 16.58% | 32.59% | 12.98% | 7.378% | -9.433% | 2.153% | -1.331% | -1.634% | -24.89% | 2.91% | -2.467% | 25.74% | 12% | -9.206% | -1.482% | -25.41% | 7.495% | -29.98% | -2.239% | -18.07% | -6.631% | -1.634% | -22%
| Z-Score (Absolute) |  | 0.11 | 0 | 0.67 | 2.01 | 0.65 | 2.22 | 1.34 | 2.6 | 1.06 | 0.62 | 0.7 | 0.21 | 0.07 | 0.09 | 1.92 | 0.27 | 0.15 | 2.06 | 0.98 | 0.68 | 0.08 | 1.96 | 0.63 | 2.32 | 0.14 | 1.38 | 0.48 | 0.09 | 1.69
| Analytical results for Cd in OREAS 234b (Certified Value 0.59 ppm) | [32 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.64 | 0.579 | < 1 | 0.4301 | 0.6404 | 0.55 | 0.66 | 0.62 | 0.53 | 0.7 | 0.6 | 0.6 | 0.61 | 0.55 | < 10 | 0.63 | < 1 | < 1 | 0.47 | 0.47 | 0.62 | 1.56 | 0.5 | 0.53113 | 0.59 | 0.55 | 0.62 | 0.56 | 0.59 |  |  | 1
| 1 | 2 | 0.64 | 0.601 | < 1 | 0.45227 | 0.63322 | 0.67 | 0.65 | 0.66 | 0.59 | 0.7 | 0.6 | 0.59 | 0.61 | 0.58 | < 10 | 0.66 | < 1 | < 1 | 0.47 | 0.51 | 0.57 | 1.59 | 0.5 | 0.59387 | 0.6 | 0.52 | 0.62 | 0.54 | 0.6 |  |  | 24
| 1 | 3 | 0.64 | 0.579 | < 1 | 0.43622 | 0.62604 | 0.65 | 0.69 | 0.65 | 0.63 | 0.7 | 0.6 | 0.62 | 0.58 | 0.61 | < 10 | 0.66 | < 1 | < 1 | 0.46 | 0.52 | 0.57 | 1.5 | 0.5 | 0.54194 | 0.6 | 0.54 | 0.64 | 0.54 | 0.6 |  |  | 16
| 1 | 4 | 0.63 | 0.573 | < 1 | 0.42876 | 0.63659 | 0.54 | 0.65 | 0.65 | 0.61 | 0.7 | 0.6 | 0.58 | 0.61 | 0.57 | < 10 | 0.64 | < 1 | < 1 | 0.49 | 0.51 | 0.59 | 1.61 | 0.5 | 0.55379 | 0.59 | 0.51 | 0.63 | 0.54 | 0.62 |  |  | 0.59468
| 1 | 5 | 0.61 | 0.585 | < 1 | 0.44899 | 0.62738 | 0.5 | 0.64 | 0.67 | 0.62 | 0.7 | 0.6 | 0.61 | 0.6 | 0.56 | < 10 | 0.66 | < 1 | < 1 | 0.49 | 0.51 | 0.59 | 1.57 | 0.5 | 0.64896 | 0.61 | 0.54 | 0.6 | 0.54 | 0.6 |  |  | 82
| 1 | 6 | 0.61 | 0.583 | < 1 | 0.4347 | 0.63647 | 0.62 | 0.67 | 0.68 | 0.58 | 0.7 | 0.6 | 0.62 | 0.57 | 0.56 | < 10 | 0.61 | < 1 | < 1 | 0.45 | 0.51 | 0.6 | 1.53 | 0.5 | 0.50247 | 0.59 | 0.49 | 0.61 | 0.53 | 0.58
| Mean |  | 0.62833 | 0.58333 |  | 0.4385 | 0.63335 | 0.58833 | 0.66 | 0.655 | 0.59333 | 0.7 | 0.6 | 0.60333 | 0.59667 | 0.57167 |  | 0.64333 |  |  | 0.47167 | 0.505 | 0.59 | 1.56 | 0.5 | 0.56203 | 0.59667 | 0.525 | 0.62 | 0.54167 | 0.59833
| Median |  | 0.635 | 0.581 |  | 0.43546 | 0.63485 | 0.585 | 0.655 | 0.655 | 0.6 | 0.7 | 0.6 | 0.605 | 0.605 | 0.565 |  | 0.65 |  |  | 0.47 | 0.51 | 0.59 | 1.565 | 0.5 | 0.54787 | 0.595 | 0.53 | 0.62 | 0.54 | 0.6
| Std Dev. |  | 0.01472 | 0.0095847 |  | 0.0098449 | 0.00564 | 0.067946 | 0.017889 | 0.020736 | 0.036148 | 1.2162e-16 | 0 | 0.01633 | 0.017512 | 0.02137 |  | 0.020656 |  |  | 0.016021 | 0.017607 | 0.018974 | 0.04 | 0 | 0.052047 | 0.008165 | 0.022583 | 0.014142 | 0.0098319 | 0.013292
| Rel.Std.Dev. |  | 2.343% | 1.643% |  | 2.245% | 0.8905% | 11.55% | 2.71% | 3.166% | 6.092% | 1.737e-14% | 0% | 2.707% | 2.935% | 3.738% |  | 3.211% |  |  | 3.397% | 3.486% | 3.216% | 2.564% | 0% | 9.261% | 1.368% | 4.302% | 2.281% | 1.815% | 2.221%
| PDM3 |  | 5.66% | -1.907% |  | -26.26% | 6.504% | -1.067% | 10.98% | 10.14% | -0.2258% | 17.71% | 0.8953% | 1.456% | 0.3347% | -3.869% |  | 8.182% |  |  | -20.69% | -15.08% | -0.7863% | 162.3% | -15.92% | -5.49% | 0.3347% | -11.72% | 4.258% | -8.914% | 0.615%
| Z-Score (Absolute) |  | 0.54 | 0.12 | 1.35 | 2.25 | 0.61 | 0.05 | 1.01 | 0.93 | 0.02 | NA | NA | 0.17 | 0.07 | 0.29 | 64.84 | 0.76 | 1.35 | 1.35 | 1.77 | 1.27 | 0.02 | 14.24 | NA | 0.44 | 0.07 | 0.98 | 0.42 | 0.74 | 0.1
| Analytical results for Ce in OREAS 234b (Certified Value 9.73 ppm) | [32 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | 9.06 | 9.897 | 10.065 | 8.3757 | 9.7619 | 10.2 | < 1 | 11 | 9.765 | 10 | 9.03 | 10 | 8.9 | 9.85 | < 10 | 9.24 | 9.72 | 10 | 10.2 | 10.4 | 9.67 | 10.22 | 9.61 | [8 empty] | 1
| 1 | 2 | 8.92 | 9.785 | 8.777 | 8.3283 | 9.6979 | 9.65 | < 1 | 11.4 | 9.367 | 11 | 8.86 | 9.96 | 9.76 | 9.56 | < 10 | 9.41 | 10.17 | 9 | 10.05 | 10.4 | 9.85 | 10.09 | 9.78 | [8 empty] | 25
| 1 | 3 | 9.13 | 9.589 | 10.026 | 8.4284 | 9.9517 | 10 | < 1 | 11.5 | 9.694 | 10 | 9.07 | 9.33 | 9.93 | 9.87 | < 10 | 9.64 | 10.03 | 9.5 | 10.55 | 10.1 | 9.7 | 10.27 | 9.66 | [8 empty] | 16
| 1 | 4 | 9.38 | 9.85 | 9.138 | 8.3011 | 9.5078 | 9.89 | < 1 | 11.3 | 9.643 | 10 | 8.95 | 9.72 | 9.59 | 9.76 | < 10 | 10.68 | 9.92 | 9.5 | 10.1 | 10.4 | 9.65 | 9.98 | 9.88 | [8 empty] | 9.7346
| 1 | 5 | 9.3 | 9.788 | 8.982 | 8.3995 | 10.2534 | 10 | < 1 | 11.3 | 9.942 | 10 | 8.93 | 9.64 | 9.34 | 9.69 | < 10 | 10.15 | 10.05 | 9 | 10.2 | 10.3 | 9.67 | 10.39 | 9.76 | [8 empty] | 83
| 1 | 6 | 9.21 | 9.79 | 9.483 | 8.248 | 10.2758 | 10.1 | < 1 | 11.6 | 9.395 | 10 | 9.19 | 9.56 | 9.43 | 9.47 | < 10 | 10.45 | 10.13 | 9.5 | 10.25 | 10.4 | 9.91 | 10.08 | 9.46
| Mean |  | 9.1667 | 9.7832 | 9.4118 | 8.3468 | 9.9081 | 9.9733 |  | 11.35 | 9.6343 | 10.1667 | 9.005 | 9.7017 | 9.4917 | 9.7 |  | 9.9283 | 10.0033 | 9.4167 | 10.225 | 10.3333 | 9.7417 | 10.1717 | 9.6917
| Median |  | 9.17 | 9.789 | 9.3105 | 8.352 | 9.8568 | 10 |  | 11.35 | 9.6685 | 10 | 8.99 | 9.68 | 9.51 | 9.725 |  | 9.895 | 10.04 | 9.5 | 10.2 | 10.4 | 9.685 | 10.155 | 9.71
| Std Dev. |  | 0.16657 | 0.10507 | 0.54252 | 0.067035 | 0.31052 | 0.18991 |  | 0.20736 | 0.22096 | 0.40825 | 0.11726 | 0.25222 | 0.36074 | 0.15975 |  | 0.58513 | 0.16367 | 0.37639 | 0.17536 | 0.12111 | 0.10998 | 0.14932 | 0.14784
| Rel.Std.Dev. |  | 1.817% | 1.074% | 5.764% | 0.8031% | 3.134% | 1.904% |  | 1.827% | 2.293% | 4.016% | 1.302% | 2.6% | 3.801% | 1.647% |  | 5.894% | 1.636% | 3.997% | 1.715% | 1.172% | 1.129% | 1.468% | 1.525%
| PDM3 |  | -5.834% | 0.4991% | -3.315% | -14.26% | 1.782% | 2.453% |  | 16.59% | -1.03% | 4.439% | -7.495% | -0.3381% | -2.495% | -0.3552% |  | 1.99% | 2.761% | -3.266% | 5.038% | 6.151% | 0.07281% | 4.49% | -0.4408%
| Z-Score (Absolute) |  | 1.26 | 0.19 | 0.68 | 3.19 | 0.49 | 0.64 | 21.65 | 3.88 | 0.16 | NA | 1.64 | 0 | 0.49 | 0 | 11.06 | 0.54 | 0.71 | 0.67 | 1.23 | 1.49 | 0.1 | 1.11 | 0.02
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 234b (Certified Value 31.4 ppm) | [32 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 30.5 | 31.08 | 27.96 | 24.3128 | 32.2183 | 32.8 | 32 | 35 | 32.1 | 29 | 31.7 | 34.6 | 32.3 | 29.9 | 24.63 | 31.6 | 29 | 34.4295 | 31.5 | 28 | 32 | 27.96 | 28 | 32.5579 | 30.7 | 35.2 | 32.2 | 31.8 | 30.3 |  |  | 1
| 1 | 2 | 30.2 | 30.97 | 27.55 | 24.2952 | 32.5323 | 32.3 | 29.9 | 34.8 | 31.8 | 29 | 32.3 | 33.7 | 31.7 | 31.6 | 26.391 | 30.6 | 29 | 34.5975 | 32.9 | 29.9 | 32 | 29.29 | 28 | 32.9914 | 30.8 | 36 | 33.3 | 31.1 | 31.1 |  |  | 26
| 1 | 3 | 30.2 | 30.96 | 28.25 | 24.1908 | 32.7554 | 32.8 | 31.9 | 34.4 | 31.8 | 28 | 31.7 | 33.7 | 30.1 | 33 | 24.03 | 31.5 | 29 | 35.3325 | 32.2 | 31.4 | 32 | 29.05 | 28 | 32.7789 | 30.7 | 35.1 | 32.1 | 31.4 | 30.7 |  |  | 16
| 1 | 4 | 30.8 | 31.03 | 27.99 | 24.1893 | 32.1989 | 31 | 31.5 | 33.9 | 32.3 | 28 | 31.7 | 33.4 | 31.5 | 31.3 | 26.509 | 31.4 | 29 | 34.9191 | 34.1 | 29.4 | 32 | 31.15 | 29 | 33.1033 | 30.7 | 35.5 | 31.7 | 30.8 | 31.3 |  |  | 31.4269
| 1 | 5 | 30.4 | 31.01 | 28.47 | 24.1636 | 32.4239 | 31.6 | 31.9 | 34.2 | 33.3 | 28 | 31 | 33.8 | 31 | 30.5 | 25.984 | 31 | 28 | 35.4925 | 34.3 | 29.4 | 32 | 26.91 | 28 | 32.9305 | 30.6 | 34.7 | 32.7 | 31.9 | 31 |  |  | 84
| 1 | 6 | 30.1 | 30.82 | 27.29 | 24.1441 | 33.0278 | 31.9 | 32.4 | 34.4 | 31.7 | 28 | 30.9 | 31.2 | 30.1 | 30.5 | 27.886 | 30.4 | 30 | 35.1703 | 34.5 | 30.8 | 32 | 28.25 | 27 | 32.7822 | 30.4 | 35.8 | 32.8 | 30.8 | 30.4
| Mean |  | 30.3667 | 30.9783 | 27.9183 | 24.216 | 32.5261 | 32.0667 | 31.6 | 34.45 | 32.1667 | 28.3333 | 31.55 | 33.4 | 31.1167 | 31.1333 | 25.905 | 31.0833 | 29 | 34.9902 | 33.25 | 29.8167 | 32 | 28.7683 | 28 | 32.8574 | 30.65 | 35.3833 | 32.4667 | 31.3 | 30.8
| Median |  | 30.3 | 30.99 | 27.975 | 24.19 | 32.4781 | 32.1 | 31.9 | 34.4 | 31.95 | 28 | 31.7 | 33.7 | 31.25 | 30.9 | 26.1875 | 31.2 | 29 | 35.0447 | 33.5 | 29.65 | 32 | 28.65 | 28 | 32.8564 | 30.7 | 35.35 | 32.45 | 31.25 | 30.85
| Std Dev. |  | 0.2582 | 0.088863 | 0.43628 | 0.070573 | 0.32144 | 0.70899 | 0.88091 | 0.39875 | 0.59889 | 0.5164 | 0.52058 | 1.1507 | 0.89088 | 1.1003 | 1.3909 | 0.49967 | 0.63246 | 0.41847 | 1.239 | 1.194 | 0 | 1.4422 | 0.63246 | 0.19251 | 0.13784 | 0.47924 | 0.57504 | 0.48166 | 0.4
| Rel.Std.Dev. |  | 0.8503% | 0.2869% | 1.563% | 0.2914% | 0.9882% | 2.211% | 2.788% | 1.157% | 1.862% | 1.823% | 1.65% | 3.445% | 2.863% | 3.534% | 5.369% | 1.608% | 2.181% | 1.196% | 3.726% | 4.005% | 0% | 5.013% | 2.259% | 0.5859% | 0.4497% | 1.354% | 1.771% | 1.539% | 1.299%
| PDM3 |  | -3.374% | -1.427% | -11.16% | -22.95% | 3.498% | 2.036% | 0.5509% | 9.62% | 2.354% | -9.844% | 0.3918% | 6.278% | -0.9871% | -0.934% | -17.57% | -1.093% | -7.722% | 11.34% | 5.801% | -5.124% | 1.824% | -8.459% | -10.9% | 4.552% | -2.472% | 12.59% | 3.309% | -0.4037% | -1.995%
| Z-Score (Absolute) |  | 0.39 | 0.08 | 1.63 | 3.5 | 0.7 | 0.47 | 0.24 | 1.68 | 0.52 | 1.42 | 0.21 | 1.15 | 0.01 | 0 | 2.64 | 0.03 | 1.08 | 1.95 | 1.07 | 0.67 | 0.44 | 1.2 | 1.58 | 0.87 | 0.24 | 2.15 | 0.67 | 0.08 | 0.17
| Analytical results for Cr in OREAS 234b (Certified Value 30 ppm) | [32 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 32 | 29.9 | 31.29 | 31.0407 | 36 | 35 | 35 | 32 | 30 | 33 | 30 | 29 | 28 | 25.569 | 30 | 26 | 31.7494 | 30 | 27 | 30 | 28.38 | 30.1971 | 25 | 29 | 28.5 | 30 | 27 | 27.4 |  |  |  | 1
| 1 | 2 | 32 | 30.3 | 31.36 | 30.7151 | 34 | 34 | 36 | 31 | 28 | 33 | 29 | 28 | 28 | 25.924 | 29 | 26 | 32.2535 | 31 | 30 | 30 | 31.38 | 30.9969 | 25 | 29 | 28.3 | 29 | 27 | 26.9 |  |  |  | 27
| 1 | 3 | 32 | 29.8 | 31.27 | 31.0082 | 35 | 33 | 36 | 33 | 28 | 32 | 30 | 29 | 29 | 24.848 | 29 | 26 | 31.2874 | 30 | 30.5 | 30 | 29.18 | 29.4438 | 25 | 29 | 28.3 | 30 | 28 | 27.3 |  |  |  | 16
| 1 | 4 | 33 | 30 | 30.92 | 30.4899 | 34 | 34 | 36 | 33 | 29 | 33 | 29 | 29 | 28 | 25.419 | 29 | 25 | 31.9169 | 35 | 29.1 | 30 | 34.14 | 29.0997 | 25 | 29 | 28.6 | 30 | 27 | 27.2 |  |  |  | 30.041
| 1 | 5 | 33 | 30.2 | 31.3 | 31.1324 | 35 | 33 | 37 | 33 | 29 | 32 | 29 | 29 | 28 | 25.317 | 29 | 25 | 32.9346 | 32 | 28.6 | 30 | 34.49 | 31.8804 | 25 | 28 | 27.8 | 29 | 28 | 27.1 |  |  |  | 85
| 1 | 6 | 32 | 30.3 | 30.76 | 31.3543 | 34 | 34 | 37 | 33 | 28 | 32 | 28 | 28 | 28 | 27.76 | 29 | 27 | 33.208 | 36 | 29.9 | 30 | 39.93 | 30.8016 | 25 | 28 | 28.1 | 30 | 28 | 26.8
| Mean |  | 32.3333 | 30.0833 | 31.15 | 30.9567 | 34.6667 | 33.8333 | 36.1667 | 32.5 | 28.6667 | 32.5 | 29.1667 | 28.6667 | 28.1667 | 25.8062 | 29.1667 | 25.8333 | 32.225 | 32.3333 | 29.1833 | 30 | 32.9167 | 30.4032 | 25 | 28.6667 | 28.2667 | 29.6667 | 27.5 | 27.1167
| Median |  | 32 | 30.1 | 31.28 | 31.0244 | 34.5 | 34 | 36 | 33 | 28.5 | 32.5 | 29 | 29 | 28 | 25.494 | 29 | 26 | 32.0852 | 31.5 | 29.5 | 30 | 32.76 | 30.4993 | 25 | 29 | 28.3 | 30 | 27.5 | 27.15
| Std Dev. |  | 0.5164 | 0.2137 | 0.24722 | 0.3082 | 0.8165 | 0.75277 | 0.75277 | 0.83666 | 0.8165 | 0.54772 | 0.75277 | 0.5164 | 0.40825 | 1.0193 | 0.40825 | 0.75277 | 0.73086 | 2.582 | 1.2671 | 0 | 4.2442 | 1.0349 | 0 | 0.5164 | 0.28752 | 0.5164 | 0.54772 | 0.23166
| Rel.Std.Dev. |  | 1.597% | 0.7104% | 0.7937% | 0.9956% | 2.355% | 2.225% | 2.081% | 2.574% | 2.848% | 1.685% | 2.581% | 1.801% | 1.449% | 3.95% | 1.4% | 2.914% | 2.268% | 7.986% | 4.342% | 0% | 12.89% | 3.404% | 0% | 1.801% | 1.017% | 1.741% | 1.992% | 0.8543%
| PDM3 |  | 7.631% | 0.141% | 3.692% | 3.048% | 15.4% | 12.62% | 20.39% | 8.186% | -4.575% | 8.186% | -2.91% | -4.575% | -6.239% | -14.1% | -2.91% | -14.01% | 7.27% | 7.631% | -2.855% | -0.1364% | 9.573% | 1.206% | -16.78% | -4.575% | -5.906% | -1.246% | -8.458% | -9.734%
| Z-Score (Absolute) |  | 0.84 | 0.08 | 0.44 | 0.38 | 1.63 | 1.35 | 2.14 | 0.9 | 0.39 | 0.9 | 0.22 | 0.39 | 0.56 | 1.36 | 0.22 | 1.35 | 0.81 | 0.84 | 0.22 | 0.06 | 1.04 | 0.19 | 1.63 | 0.39 | 0.53 | 0.06 | 0.79 | 0.92
| Analytical results for Cs in OREAS 234b (Certified Value 0.71 ppm) | [32 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.78 | 0.7 | 0.61512 | 0.64 | 0.67 | 0.72 | 0.71 | 0.8 | 0.69 | 0.69 | 0.65 | 0.71 | 0.61 | 0.5 | 0.71 | 0.65 | 0.75 | 0.74 | 0.7 | [12 empty] | 1
| 1 | 2 | 0.79 | 0.7 | 0.59984 | 0.55 | 0.68 | 0.76 | 0.71 | 0.8 | 0.7 | 0.73 | 0.72 | 0.71 | 0.62 | 0.5 | 0.71 | 0.63 | 0.77 | 0.75 | 0.72 | [12 empty] | 28
| 1 | 3 | 0.78 | 0.72 | 0.59514 | 0.59 | 0.68 | 0.75 | 0.71 | 0.8 | 0.68 | 0.72 | 0.75 | 0.73 | 0.68 | 0.5 | 0.73 | 0.64 | 0.76 | 0.74 | 0.72 | [12 empty] | 16
| 1 | 4 | 0.82 | 0.7 | 0.58421 | 0.61 | 0.68 | 0.77 | 0.72 | 0.8 | 0.69 | 0.75 | 0.72 | 0.72 | 0.7 | 0.5 | 0.72 | 0.64 | 0.75 | 0.73 | 0.74 | [12 empty] | 0.71487
| 1 | 5 | 0.79 | 0.71 | 0.59464 | 0.58 | 0.67 | 0.75 | 0.72 | 0.8 | 0.68 | 0.71 | 0.7 | 0.7 | 0.73 | 0.5 | 0.71 | 0.59 | 0.76 | 0.74 | 0.72 | [12 empty] | 86
| 1 | 6 | 0.77 | 0.72 | 0.57663 | 0.62 | 0.66 | 0.77 | 0.68 | 0.8 | 0.69 | 0.71 | 0.72 | 0.69 | 0.72 | 0.5 | 0.71 | 0.65 | 0.74 | 0.73 | 0.72
| Mean |  | 0.78833 | 0.70833 | 0.59426 | 0.59833 | 0.67333 | 0.75333 | 0.70833 | 0.8 | 0.68833 | 0.71833 | 0.71 | 0.71 | 0.67667 | 0.5 | 0.715 | 0.63333 | 0.755 | 0.73833 | 0.72
| Median |  | 0.785 | 0.705 | 0.59489 | 0.6 | 0.675 | 0.755 | 0.71 | 0.8 | 0.69 | 0.715 | 0.72 | 0.71 | 0.69 | 0.5 | 0.71 | 0.64 | 0.755 | 0.74 | 0.72
| Std Dev. |  | 0.017224 | 0.0098319 | 0.01326 | 0.031885 | 0.008165 | 0.018619 | 0.01472 | 1.2162e-16 | 0.0075277 | 0.020412 | 0.033466 | 0.014142 | 0.050859 | 0 | 0.0083666 | 0.022509 | 0.010488 | 0.0075277 | 0.012649
| Rel.Std.Dev. |  | 2.185% | 1.388% | 2.231% | 5.329% | 1.213% | 2.472% | 2.078% | 1.52e-14% | 1.094% | 2.842% | 4.714% | 1.992% | 7.516% | 0% | 1.17% | 3.554% | 1.389% | 1.02% | 1.757%
| PDM3 |  | 10.28% | -0.9139% | -16.87% | -16.3% | -5.81% | 5.381% | -0.9139% | 11.91% | -3.712% | 0.4849% | -0.6808% | -0.6808% | -5.344% | -30.06% | 0.01865% | -11.41% | 5.614% | 3.283% | 0.7181%
| Z-Score (Absolute) |  | 1.86 | 0.04 | 2.75 | 2.66 | 0.87 | 1.03 | 0.04 | NA | 0.52 | 0.2 | 0 | 0 | 0.79 | NA | 0.12 | 1.82 | 1.07 | 0.67 | 0.24
Results from laboratories 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 234b (Certified Value 160 ppm) | [32 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 179 | 164.59 | 145.73 | 179.71 | 167.0629 | 169 | 168 | 152 | 163.3 | 160 | 151.8 | 167 | 163.5 | 161 | 159.41 | 165 | 171 | 160.7369 | 159 | 136.6 | 164 | 159.09 | 156.254 | 158 | 160.0618 | 164 | 141.162 | 167 | 165.6 | 154.4 | 158 | 1
| 1 | 2 | 179.4 | 164.92 | 147.23 | 179.62 | 166.2899 | 153 | 166 | 155 | 156 | 157 | 157.1 | 166 | 160 | 165.5 | 158.986 | 160.5 | 170 | 160.6807 | 156 | 149.09 | 158 | 160.65 | 155.5648 | 157 | 162.2882 | 163 | 138.98 | 170 | 162.8 | 156 | 160 | 29
| 1 | 3 | 180.4 | 165.74 | 152.02 | 180.13 | 167.6078 | 161 | 164 | 154 | 159.9 | 155 | 155.3 | 166.5 | 160.5 | 169.5 | 154.774 | 162 | 173 | 164.7094 | 159 | 152.41 | 160 | 159.52 | 150.6776 | 158 | 164.2552 | 160.5 | 144.542 | 162 | 167.1 | 157.4 | 160 | 16
| 1 | 4 | 182.1 | 164.27 | 150.64 | 177.04 | 164.0469 | 153 | 165 | 153 | 163.9 | 158 | 152.1 | 165 | 164 | 167 | 158.404 | 162.5 | 170 | 166.6239 | 164 | 150.1 | 163 | 162.68 | 154.6336 | 159 | 168.8216 | 162.5 | 142.42 | 167 | 163.2 | 152.6 | 161 | 160.3635
| 1 | 5 | 180.1 | 164.94 | 149.28 | 179.96 | 168.7603 | 148 | 164 | 155 | 163.4 | 156 | 150.1 | 165.5 | 160.5 | 162.5 | 162.056 | 160 | 165 | 167.8675 | 167 | 146.5 | 163 | 159.79 | 151.0184 | 157 | 166.8727 | 162 | 137.683 | 165 | 162 | 157.9 | 161 | 87
| 1 | 6 | 176.9 | 165.3 | 146.98 | 182.48 | 168.8485 | 159 | 168 | 155 | 157.3 | 156 | 148.7 | 166.5 | 160.5 | 160.5 | 167.19 | 159 | 174 | 164.2916 | 162 | 150.15 | 162 | 161.28 | 153.5832 | 153 | 162.6486 | 161.5 | 154.362 | 163 | 162.4 | 153.2 | 156
| Mean |  | 179.65 | 164.96 | 148.6467 | 179.8233 | 167.1027 | 157.1667 | 165.8333 | 154 | 160.6333 | 157 | 152.5167 | 166.0833 | 161.5 | 164.3333 | 160.1367 | 161.5 | 170.5 | 164.1516 | 161.1667 | 147.475 | 161.6667 | 160.5017 | 153.6219 | 157 | 164.158 | 162.25 | 143.1915 | 165.6667 | 163.85 | 155.25 | 159.3333
| Median |  | 179.75 | 164.93 | 148.255 | 179.835 | 167.3353 | 156 | 165.5 | 154.5 | 161.6 | 156.5 | 151.95 | 166.25 | 160.5 | 164 | 159.198 | 161.25 | 170.5 | 164.5005 | 160.5 | 149.595 | 162.5 | 160.22 | 154.1084 | 157.5 | 163.4519 | 162.25 | 141.791 | 166 | 163 | 155.2 | 160
| Std Dev. |  | 1.7213 | 0.51749 | 2.4102 | 1.7304 | 1.7917 | 7.4409 | 1.8348 | 1.2649 | 3.4221 | 1.7889 | 3.1575 | 0.73598 | 1.7607 | 3.587 | 4.1724 | 2.1448 | 3.1464 | 2.9651 | 3.9707 | 5.6608 | 2.2509 | 1.3294 | 2.3313 | 2.0976 | 3.215 | 1.2145 | 5.99 | 2.9439 | 2.0335 | 2.198 | 1.9664
| Rel.Std.Dev. |  | 0.9582% | 0.3137% | 1.621% | 0.9623% | 1.072% | 4.734% | 1.106% | 0.8214% | 2.13% | 1.139% | 2.07% | 0.4431% | 1.09% | 2.183% | 2.606% | 1.328% | 1.845% | 1.806% | 2.464% | 3.839% | 1.392% | 0.8282% | 1.518% | 1.336% | 1.958% | 0.7485% | 4.183% | 1.777% | 1.241% | 1.416% | 1.234%
| PDM3 |  | 12.03% | 2.866% | -7.306% | 12.13% | 4.202% | -1.994% | 3.411% | -3.968% | 0.1682% | -2.097% | -4.893% | 3.567% | 0.7087% | 2.475% | -0.1415% | 0.7087% | 6.321% | 2.362% | 0.5008% | -8.037% | 0.8126% | 0.08613% | -4.204% | -2.097% | 2.366% | 1.176% | -10.71% | 3.307% | 2.174% | -3.189% | -0.6424%
| Z-Score (Absolute) |  | 2.82 | 0.54 | 2 | 2.85 | 0.87 | 0.67 | 0.67 | 1.17 | 0.13 | 0.7 | 1.4 | 0.71 | 0 | 0.44 | 0.21 | 0 | 1.4 | 0.41 | 0.05 | 2.18 | 0.03 | 0.16 | 1.23 | 0.7 | 0.41 | 0.12 | 2.85 | 0.65 | 0.37 | 0.97 | 0.34
| Analytical results for Dy in OREAS 234b (Certified Value 2.05 ppm) | [32 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 2.097 | 1.898 | 1.7939 | 2.3 | 2.58 | 1.9 | [25 empty] | 1
| 1 | 2 | 2.086 | 1.861 | 1.7784 | 2.1 | 2.5 | 1.7 | [25 empty] | 30
| 1 | 3 | 2.058 | 2.1 | 1.7724 | 2.3 | 2.49 | 1.6 | [25 empty] | 16
| 1 | 4 | 2.1 | 1.961 | 1.7334 | 2.3 | 2.56 | 1.8 | [25 empty] | 2.0461
| 1 | 5 | 2.109 | 1.965 | 1.7726 | 2.2 | 2.36 | 1.7 | [25 empty] | 88
| 1 | 6 | 2.097 | 1.997 | 1.6966 | 2.3 | 2.42 | 1.6
| Mean |  | 2.0912 | 1.9637 | 1.7579 | 2.25 | 2.485 | 1.7167
| Median |  | 2.097 | 1.963 | 1.7725 | 2.3 | 2.495 | 1.7
| Std Dev. |  | 0.017837 | 0.083138 | 0.036086 | 0.083666 | 0.083367 | 0.1169
| Rel.Std.Dev. |  | 0.853% | 4.234% | 2.053% | 3.718% | 3.355% | 6.81%
| PDM3 |  | 2.202% | -4.029% | -14.09% | 9.965% | 21.45% | -16.1%
| Z-Score (Absolute) |  | 0.17 | 0.17 | 0.74 | 0.61 | 1.25 | 0.85
| Analytical results for Er in OREAS 234b (Certified Value 1.18 ppm) | [32 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.204 | 1.077 | 1.0965 | 1.4 | 1.47 | 1 | [25 empty] | 1
| 1 | 2 | 1.216 | 1.064 | 1.097 | 1.3 | 1.44 | 0.9 | [25 empty] | 31
| 1 | 3 | 1.218 | 1.196 | 1.087 | 1.3 | 1.41 | 0.9 | [25 empty] | 16
| 1 | 4 | 1.243 | 1.139 | 1.0656 | 1.3 | 1.44 | 0.95 | [25 empty] | 1.1783
| 1 | 5 | 1.241 | 1.128 | 1.0724 | 1.3 | 1.34 | 0.85 | [25 empty] | 89
| 1 | 6 | 1.216 | 1.147 | 1.0812 | 1.3 | 1.38 | 0.85
| Mean |  | 1.223 | 1.1252 | 1.0833 | 1.3167 | 1.4133 | 0.90833
| Median |  | 1.217 | 1.1335 | 1.0841 | 1.3 | 1.425 | 0.9
| Std Dev. |  | 0.015543 | 0.048487 | 0.012763 | 0.040825 | 0.047188 | 0.058452
| Rel.Std.Dev. |  | 1.271% | 4.309% | 1.178% | 3.101% | 3.339% | 6.435%
| PDM3 |  | 3.794% | -4.509% | -8.064% | 11.74% | 19.95% | -22.91%
| Z-Score (Absolute) |  | 0.28 | 0.28 | 0.52 | 0.82 | 1.38 | 1.54
| Analytical results for Eu in OREAS 234b (Certified Value 0.45 ppm) | [32 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 0.454 | 0.361 | 0.5125 | 0.6 | 0.55 | 0.4 | [25 empty] | 1
| 1 | 2 | 0.456 | 0.361 | 0.50992 | 0.5 | 0.53 | 0.3 | [25 empty] | 32
| 1 | 3 | 0.446 | 0.406 | 0.49427 | 0.5 | 0.54 | 0.3 | [25 empty] | 16
| 1 | 4 | 0.449 | 0.399 | 0.47807 | 0.5 | 0.56 | 0.3 | [25 empty] | 0.44925
| 1 | 5 | 0.46 | 0.383 | 0.48973 | 0.5 | 0.49 | 0.3 | [25 empty] | 90
| 1 | 6 | 0.466 | 0.382 | 0.47535 | 0.5 | 0.52 | 0.3
| Mean |  | 0.45517 | 0.382 | 0.49331 | 0.51667 | 0.53167 | 0.31667
| Median |  | 0.455 | 0.3825 | 0.492 | 0.5 | 0.535 | 0.3
| Std Dev. |  | 0.0072778 | 0.018698 | 0.015573 | 0.040825 | 0.024833 | 0.040825
| Rel.Std.Dev. |  | 1.599% | 4.895% | 3.157% | 7.902% | 4.671% | 12.89%
| PDM3 |  | 1.318% | -14.97% | 9.808% | 15.01% | 18.35% | -29.51%
| Z-Score (Absolute) |  | 0.26 | 1.25 | 0.26 | 0.57 | 0.78 | 2.13
| Analytical results for Fe in OREAS 234b (Certified Value 5.71 wt.%) | [32 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 5.91 | 5.678 | 3.8978 | 5.5312 | 5.8819 | 6.04 | 6.022 | 6.24 | 5.67 | 5.47 | 5.47 | 5.74 | 5.9 | 5.71 | 5.6419 | 6.01 | 5.75 | > 5 | 6.05 | 4.67 | 5.69 | 6.39 | 5.9404 | 5.141 | 5.54 | 5.57 | 5.72 | 5.66 | 5.43 |  |  | 1
| 1 | 2 | 5.91 | 5.631 | 3.8811 | 5.5459 | 5.8277 | 5.92 | 5.973 | 6.29 | 5.6 | 5.44 | 5.57 | 5.76 | 5.82 | 5.79 | 5.5694 | 5.75 | 5.75 | > 5 | 6.05 | 5.11 | 5.38 | 6.05 | 5.8971 | 5.084 | 5.51 | 5.7 | 5.64 | 5.62 | 5.56 |  |  | #N/A
| 1 | 3 | 5.9 | 5.675 | 3.9386 | 5.532 | 5.8746 | 6.05 | 5.896 | 6.23 | 5.54 | 5.48 | 5.5 | 5.68 | 5.41 | 5.79 | 5.5646 | 5.82 | 5.85 | > 5 | 6.05 | 5.2 | 5.66 | 6.83 | 5.7594 | 5.178 | 5.54 | 5.5 | 5.65 | 5.66 | 5.56 |  |  | 16
| 1 | 4 | 5.99 | 5.698 | 3.8735 | 5.4919 | 5.851 | 5.82 | 5.868 | 6.14 | 5.72 | 5.49 | 5.48 | 5.64 | 5.75 | 5.78 | 5.6185 | 5.83 | 5.61 | > 5 | 6.07 | 4.99 | 5.88 | 6.03 | 5.7619 | 5.2125 | 5.54 | 5.68 | 5.7 | 5.62 | 5.55 |  |  | 5.7077
| 1 | 5 | 5.97 | 5.754 | 3.9313 | 5.5103 | 5.9233 | 5.71 | 5.791 | 6.29 | 5.88 | 5.51 | 5.5 | 5.7 | 5.67 | 5.68 | 5.705 | 5.83 | 5.55 | > 5 | 6.05 | 4.93 | 5.64 | 6.01 | 5.8825 | 5.0905 | 5.47 | 5.46 | 5.67 | 5.81 | 5.63 |  |  | 91
| 1 | 6 | 5.84 | 5.701 | 3.8226 | 5.5163 | 5.9206 | 5.87 | 5.959 | 6.18 | 5.6 | 5.49 | 5.49 | 5.57 | 5.69 | 5.68 | 5.8341 | 5.8 | 5.92 | > 5 | 6.07 | 5.16 | 5.7 | 6.7 | 5.8255 | 5.0215 | 5.47 | 5.48 | 5.66 | 5.66 | 5.59
| Mean |  | 5.92 | 5.6895 | 3.8908 | 5.5213 | 5.8799 | 5.9017 | 5.9182 | 6.2283 | 5.6683 | 5.48 | 5.5017 | 5.6817 | 5.7067 | 5.7383 | 5.6556 | 5.84 | 5.7383 |  | 6.0567 | 5.01 | 5.6583 | 6.335 | 5.8445 | 5.1212 | 5.5117 | 5.565 | 5.6733 | 5.6717 | 5.5533
| Median |  | 5.91 | 5.688 | 3.8895 | 5.5237 | 5.8783 | 5.895 | 5.9275 | 6.235 | 5.635 | 5.485 | 5.495 | 5.69 | 5.72 | 5.745 | 5.6302 | 5.825 | 5.75 |  | 6.05 | 5.05 | 5.675 | 6.22 | 5.854 | 5.1158 | 5.525 | 5.535 | 5.665 | 5.66 | 5.56
| Std Dev. |  | 0.053666 | 0.040322 | 0.042463 | 0.019121 | 0.037766 | 0.13106 | 0.083228 | 0.059805 | 0.12106 | 0.023664 | 0.035449 | 0.069402 | 0.16836 | 0.054191 | 0.10158 | 0.088544 | 0.13977 |  | 0.010328 | 0.19545 | 0.16105 | 0.36363 | 0.074613 | 0.069599 | 0.034303 | 0.10387 | 0.030768 | 0.070545 | 0.067132
| Rel.Std.Dev. |  | 0.9065% | 0.7087% | 1.091% | 0.3463% | 0.6423% | 2.221% | 1.406% | 0.9602% | 2.136% | 0.4318% | 0.6443% | 1.222% | 2.95% | 0.9444% | 1.796% | 1.516% | 2.436% |  | 0.1705% | 3.901% | 2.846% | 5.74% | 1.277% | 1.359% | 0.6224% | 1.867% | 0.5423% | 1.244% | 1.209%
| PDM3 |  | 3.72% | -0.3181% | -31.83% | -3.266% | 3.017% | 3.399% | 3.688% | 9.122% | -0.689% | -3.989% | -3.609% | -0.4554% | -0.01735% | 0.5375% | -0.9123% | 2.319% | 0.5375% |  | 6.115% | -12.22% | -0.8642% | 10.99% | 2.397% | -10.27% | -3.434% | -2.499% | -0.6014% | -0.6306% | -2.704%
| Z-Score (Absolute) |  | 1 | 0.05 | 7.34 | 0.64 | 0.83 | 0.92 | 0.99 | 2.26 | 0.04 | 0.81 | 0.72 | 0.02 | 0.12 | 0.25 | 0.09 | 0.67 | 0.25 | NA | 1.56 | 2.74 | 0.08 | 2.7 | 0.69 | 2.28 | 0.68 | 0.46 | 0.02 | 0.02 | 0.51
| Analytical results for Ga in OREAS 234b (Certified Value 9.78 ppm) | [32 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 9 | 10.4 | 9.3022 | 8.6541 | 9.96 | 11.5 | 10.6 | 9.89 | 10 | 10 | 9.8 | 10.4 | 9.05 | 10.3 | 6 | 10.2 | 8.2 | 11.1 | 12.18 | 10 | 9.74 | 11.2 | 9.9 | 7.9 | 8.58 | [6 empty] | 1
| 1 | 2 | 9.1 | 10.3 | 9.3168 | 8.6797 | 9.54 | 11.2 | 10.4 | 9.65 | 10 | 10 | 9.72 | 10.25 | 9.49 | 10 | 6 | 10.3 | 9.1 | 10.9 | 12.23 | 9 | 9.69 | 11.3 | 10.2 | 7.66 | 8.59 | [6 empty] | 33
| 1 | 3 | 9 | 10.2 | 9.2598 | 8.5823 | 9.63 | 11.2 | 10.2 | 9.8 | 10 | 9 | 9.61 | 9.48 | 10 | 10.2 | 6 | 10.2 | 9.1 | 11.3 | 12 | 9 | 10.1 | 11.8 | 10 | 7.89 | 8.64 | [6 empty] | 16
| 1 | 4 | 9.4 | 10.4 | 9.2429 | 8.4535 | 9.49 | 11.6 | 10.2 | 9.98 | 10 | 10 | 9.62 | 10.05 | 9.55 | 10.15 | 6 | 11.1 | 8.6 | 11.2 | 12.99 | 10 | 9.77 | 11.4 | 9.8 | 7.84 | 8.84 | [6 empty] | 9.7786
| 1 | 5 | 9.3 | 10.3 | 9.2676 | 8.6286 | 9.48 | 10.7 | 10.4 | 10.18 | 10 | 9 | 9.7 | 10 | 9.27 | 10.2 | 6 | 11.4 | 8.5 | 10.7 | 11.95 | 9 | 9.89 | 11.3 | 10 | 7.89 | 8.73 | [6 empty] | 92
| 1 | 6 | 9.3 | 10.3 | 9.3199 | 8.657 | 9.46 | 10.7 | 10.5 | 9.59 | 10 | 9 | 9.78 | 9.73 | 9.29 | 9.84 | 7 | 10.9 | 8.8 | 11.3 | 12.74 | 9 | 9.74 | 11.3 | 10.2 | 7.77 | 8.45
| Mean |  | 9.1833 | 10.3167 | 9.2849 | 8.6092 | 9.5933 | 11.15 | 10.3833 | 9.8483 | 10 | 9.5 | 9.705 | 9.985 | 9.4417 | 10.115 | 6.1667 | 10.6833 | 8.7167 | 11.0833 | 12.3483 | 9.3333 | 9.8217 | 11.3833 | 10.0167 | 7.825 | 8.6383
| Median |  | 9.2 | 10.3 | 9.2849 | 8.6413 | 9.515 | 11.2 | 10.4 | 9.845 | 10 | 9.5 | 9.71 | 10.025 | 9.39 | 10.175 | 6 | 10.6 | 8.7 | 11.15 | 12.205 | 9 | 9.755 | 11.3 | 10 | 7.865 | 8.615
| Std Dev. |  | 0.17224 | 0.075277 | 0.032357 | 0.083188 | 0.1897 | 0.38341 | 0.16021 | 0.21794 | 0 | 0.54772 | 0.07893 | 0.33673 | 0.32609 | 0.16658 | 0.40825 | 0.51929 | 0.35449 | 0.24014 | 0.4213 | 0.5164 | 0.15198 | 0.2137 | 0.16021 | 0.094393 | 0.13438
| Rel.Std.Dev. |  | 1.876% | 0.7297% | 0.3485% | 0.9663% | 1.977% | 3.439% | 1.543% | 2.213% | 0% | 5.766% | 0.8133% | 3.372% | 3.454% | 1.647% | 6.62% | 4.861% | 4.067% | 2.167% | 3.412% | 5.533% | 1.547% | 1.877% | 1.599% | 1.206% | 1.556%
| PDM3 |  | -6.087% | 5.503% | -5.049% | -11.96% | -1.894% | 14.02% | 6.184% | 0.7133% | 2.264% | -2.849% | -0.7525% | 2.111% | -3.445% | 3.44% | -36.94% | 9.252% | -10.86% | 13.34% | 26.28% | -4.553% | 0.4405% | 16.41% | 2.435% | -19.98% | -11.66%
| Z-Score (Absolute) |  | 0.8 | 0.56 | 0.67 | 1.48 | 0.31 | 1.56 | 0.64 | 0 | NA | NA | 0.17 | 0.16 | 0.49 | 0.32 | NA | 1 | 1.35 | 1.48 | 2.99 | NA | 0.03 | 1.84 | 0.2 | 2.42 | 1.45
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 234b (Certified Value 1.81 ppm) | [32 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.956 | 1.513 | 1.959 | 2 | 1.89 | 1.7 | [25 empty] | 1
| 1 | 2 | 1.952 | 1.466 | 1.9259 | 1.9 | 1.85 | 1.6 | [25 empty] | 34
| 1 | 3 | 1.935 | 1.659 | 1.8884 | 2 | 1.88 | 1.5 | [25 empty] | 16
| 1 | 4 | 1.948 | 1.552 | 1.8695 | 2 | 1.9 | 1.7 | [25 empty] | 1.8084
| 1 | 5 | 1.945 | 1.529 | 1.9085 | 2.1 | 1.76 | 1.6 | [25 empty] | 93
| 1 | 6 | 1.959 | 1.575 | 1.9023 | 2 | 1.78 | 1.5
| Mean |  | 1.9492 | 1.549 | 1.9089 | 2 | 1.8433 | 1.6
| Median |  | 1.95 | 1.5405 | 1.9054 | 2 | 1.865 | 1.6
| Std Dev. |  | 0.008612 | 0.065376 | 0.03102 | 0.063246 | 0.059554 | 0.089443
| Rel.Std.Dev. |  | 0.4418% | 4.221% | 1.625% | 3.162% | 3.231% | 5.59%
| PDM3 |  | 7.784% | -14.34% | 5.558% | 10.59% | 1.932% | -11.52%
| Z-Score (Absolute) |  | 0.5 | 2.24 | 0.22 | 0.85 | 0.22 | 1.89
| Analytical results for Ge in OREAS 234b (Certified Value 0.1 ppm) | [32 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 2
| 1 | 1 | < 0.1 | 0.08 | 0.232 | < 0.2 | < 0.1 | 0.11 | 2 | 0.12 | 0.06 | 0.09 | 0.14 | < 10 | 0.1 | 0.15 | 0.15 | 0.2 | 0.1 | 0.1 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.08 | 0.23721 | < 0.2 | < 0.1 | 0.12 | 2.2 | 0.12 | 0.05 | 0.09 | 0.13 | < 10 | < 0.1 | 0.14 | 0.15 | 0.2 | 0.11 | 0.1 | [13 empty] | 35
| 1 | 3 | < 0.1 | 0.08 | 0.22091 | < 0.2 | < 0.1 | 0.09 | 2 | 0.13 | 0.05 | 0.1 | 0.13 | < 10 | 0.1 | 0.15 | 0.1 | 0.2 | 0.1 | 0.09 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.08 | 0.22185 | < 0.2 | < 0.1 | 0.09 | 2 | 0.11 | 0.05 | 0.1 | 0.14 | < 10 | 0.1 | 0.15 | 0.11 | 0.2 | 0.09 | 0.1 | [13 empty] | 0.104
| 1 | 5 | < 0.1 | 0.08 | 0.23419 | < 0.2 | < 0.1 | 0.06 | 2 | 0.11 | 0.06 | 0.11 | 0.14 | < 10 | < 0.1 | 0.13 | 0.09 | 0.2 | 0.1 | 0.11 | [13 empty] | 94
| 1 | 6 | < 0.1 | 0.08 | 0.22379 | < 0.2 | < 0.1 | 0.09 | 2 | 0.15 | 0.05 | 0.09 | 0.13 | < 10 | < 0.1 | 0.13 | 0.1 | 0.2 | 0.11 | 0.09
| Mean |  |  | 0.08 | 0.22833 |  |  | 0.093333 | 2.0333 | 0.12333 | 0.053333 | 0.096667 | 0.135 |  | 0.1 | 0.14167 | 0.11667 | 0.2 | 0.10167 | 0.098333
| Median |  |  | 0.08 | 0.22789 |  |  | 0.09 | 2 | 0.12 | 0.05 | 0.095 | 0.135 |  | 0.1 | 0.145 | 0.105 | 0.2 | 0.1 | 0.1
| Std Dev. |  |  | 0 | 0.0069898 |  |  | 0.020656 | 0.08165 | 0.015055 | 0.005164 | 0.008165 | 0.0054772 |  | 1.6997e-17 | 0.0098319 | 0.026583 | 3.0405e-17 | 0.0075277 | 0.0075277
| Rel.Std.Dev. |  |  | 0% | 3.061% |  |  | 22.13% | 4.016% | 12.21% | 9.682% | 8.447% | 4.057% |  | 1.7e-14% | 6.94% | 22.79% | 1.52e-14% | 7.404% | 7.655%
| PDM3 |  |  | -23.08% | 119.5% |  |  | -10.26% | 1855% | 18.59% | -48.72% | -7.051% | 29.81% |  | -3.846% | 36.22% | 12.18% | 92.31% | -2.244% | -5.449%
| Z-Score (Absolute) |  | 1.21 | 0.5 | 3.03 | 0.02 | 1.21 | 0.18 | NA | 0.54 | 1.13 | 0.1 | 0.81 | 116.6 | NA | 0.97 | 0.38 | 2.36 | 0.02 | 0.06
Results from laboratories 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 234b (Certified Value 0.41 ppm) | [32 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 2
| 1 | 1 | 0.35 | 0.42 | 0.38135 | 0.4 | 0.33 | 0.45 | 0.57 | 0.3 | 0.4 | 0.47 | 0.43 | 0.38 | 0.48 | < 0.5 | 0.44 | 0.57 | 0.42 | 0.47 | 0.37 | [12 empty] | 1
| 1 | 2 | 0.37 | 0.44 | 0.37598 | 0.3 | 0.24 | 0.44 | 0.55 | 0.3 | 0.44 | 0.42 | 0.47 | 0.41 | 0.4 | < 0.5 | 0.45 | 0.62 | 0.44 | 0.49 | 0.35 | [12 empty] | 19
| 1 | 3 | 0.31 | 0.43 | 0.36512 | 0.4 | 0.35 | 0.44 | 0.56 | 0.3 | 0.41 | 0.48 | 0.45 | 0.41 | 0.43 | < 0.5 | 0.45 | 0.63 | 0.42 | 0.48 | 0.36 | [12 empty] | 16
| 1 | 4 | 0.32 | 0.43 | 0.35649 | 0.3 | 0.31 | 0.41 | 0.57 | 0.3 | 0.44 | 0.44 | 0.4 | 0.4 | 0.49 | < 0.5 | 0.44 | 0.59 | 0.4 | 0.49 | 0.37 | [12 empty] | 0.4091
| 1 | 5 | 0.32 | 0.44 | 0.35604 | 0.4 | 0.39 | 0.39 | 0.55 | 0.3 | 0.42 | 0.44 | 0.44 | 0.41 | 0.4 | < 0.5 | 0.43 | 0.58 | 0.43 | 0.51 | 0.38 | [12 empty] | 95
| 1 | 6 | 0.33 | 0.44 | 0.35559 | 0.4 | 0.36 | 0.38 | 0.53 | 0.3 | 0.46 | 0.42 | 0.47 | 0.39 | 0.43 | < 0.5 | 0.45 | 0.57 | 0.42 | 0.48 | 0.36
| Mean |  | 0.33333 | 0.43333 | 0.3651 | 0.36667 | 0.33 | 0.41833 | 0.555 | 0.3 | 0.42833 | 0.445 | 0.44333 | 0.4 | 0.43833 |  | 0.44333 | 0.59333 | 0.42167 | 0.48667 | 0.365
| Median |  | 0.325 | 0.435 | 0.3608 | 0.4 | 0.34 | 0.425 | 0.555 | 0.3 | 0.43 | 0.44 | 0.445 | 0.405 | 0.43 |  | 0.445 | 0.585 | 0.42 | 0.485 | 0.365
| Std Dev. |  | 0.022509 | 0.008165 | 0.011218 | 0.05164 | 0.051769 | 0.029269 | 0.015166 | 0 | 0.022286 | 0.0251 | 0.026583 | 0.012649 | 0.038687 |  | 0.008165 | 0.02582 | 0.013292 | 0.013663 | 0.010488
| Rel.Std.Dev. |  | 6.753% | 1.884% | 3.073% | 14.08% | 15.69% | 6.997% | 2.733% | 0% | 5.203% | 5.64% | 5.996% | 3.162% | 8.826% |  | 1.842% | 4.352% | 3.152% | 2.807% | 2.873%
| PDM3 |  | -18.52% | 5.925% | -10.76% | -10.37% | -19.33% | 2.258% | 35.67% | -26.67% | 4.703% | 8.777% | 8.369% | -2.223% | 7.147% |  | 8.369% | 45.04% | 3.073% | 18.96% | -10.78%
| Z-Score (Absolute) |  | 1.48 | 0.13 | 0.97 | 0.94 | 1.54 | 0.11 | 2.1 | NA | 0.05 | 0.32 | 0.3 | 0.4 | 0.22 | 2.83 | 0.3 | 2.72 | 0.05 | 1 | 0.97
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 234b (Indicative Value 0.048 ppm) | [32 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | < 0.01 | 0.02 | < 1 | 0.09 | < 1 | < 0.1 | < 5 | 0.643 | 0.04 | 0.01 | < 0.01 | 0.03 | < 1 | 0.09 | 0.012 | 0.06 | 0.95 | 0.05 | < 0.08 | 0.04 | 0.124 | 0.31 | < 0.01 | 0.024 | 0.38 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.02 | < 1 | 0.1 | < 1 | < 0.1 | < 5 | 0.583 | 0.02 | 0.03 | < 0.01 | 0.03 | < 1 | 0.12 | 0.01 | 0.07 | 0.2 | 0.05 | < 0.08 | 0.04 | 0.118 | 0.32 | < 0.01 | 0.022 | 0.41 | [6 empty] | 4
| 1 | 3 | < 0.01 | 0.02 | < 1 | 0.07 | < 1 | < 0.1 | < 5 | 0.573 | 0.01 | 0.03 | < 0.01 | 0.02 | < 1 | 0.08 | 0.013 | 0.07 | < 0.1 | 0.05 | < 0.08 | 0.04 | 0.102 | 0.28 | < 0.01 | 0.023 | 0.38 | [6 empty] | 16
| 1 | 4 | < 0.01 | 0.02 | < 1 | 0.09 | < 1 | < 0.1 | < 5 | 0.567 | < 0.01 | 0.02 | < 0.01 | 0.04 | < 1 | 0.11 | 0.011 | 0.05 | 0.6 | 0.05 | < 0.08 | 0.03 | 0.125 | 0.28 | < 0.01 | 0.022 | 0.37 | [6 empty] | 0.048114
| 1 | 5 | < 0.01 | 0.02 | < 1 | 0.07 | < 1 | < 0.1 | < 5 | 0.552 | < 0.01 | 0.03 | < 0.01 | 0.02 | < 1 | 0.12 | 0.016 | 0.06 | 0.44 | 0.05 | < 0.08 | 0.04 | 0.119 | 0.28 | < 0.01 | 0.02 | 0.38 | [6 empty] | 12
| 1 | 6 | < 0.01 | 0.02 | < 1 | 0.07 | < 1 | < 0.1 | < 5 | 0.558 | 0.02 | 0.02 | < 0.01 | 0.04 | < 1 | 0.08 | 0.011 | 0.03 | < 0.1 | 0.05 | < 0.08 | 0.03 | 0.125 | 0.29 | < 0.01 | 0.022 | 0.37
| Mean |  |  | 0.02 |  | 0.081667 |  |  |  | 0.57933 | 0.0225 | 0.023333 |  | 0.03 |  | 0.1 | 0.012167 | 0.056667 | 0.5475 | 0.05 |  | 0.036667 | 0.11883 | 0.29333 |  | 0.022167 | 0.38167
| Median |  |  | 0.02 |  | 0.08 |  |  |  | 0.57 | 0.02 | 0.025 |  | 0.03 |  | 0.1 | 0.0115 | 0.06 | 0.52 | 0.05 |  | 0.04 | 0.1215 | 0.285 |  | 0.022 | 0.38
| Std Dev. |  |  | 0 |  | 0.013292 |  |  |  | 0.033049 | 0.012583 | 0.008165 |  | 0.0089443 |  | 0.018974 | 0.002137 | 0.015055 | 0.31468 | 7.6012e-18 |  | 0.005164 | 0.0087958 | 0.017512 |  | 0.0013292 | 0.01472
| Rel.Std.Dev. |  |  | 0% |  | 16.28% |  |  |  | 5.705% | 55.92% | 34.99% |  | 29.81% |  | 18.97% | 17.56% | 26.57% | 57.48% | 1.52e-14% |  | 14.08% | 7.402% | 5.97% |  | 5.996% | 3.857%
| PDM3 |  |  | -58.43% |  | 69.74% |  |  |  | 1104% | -53.24% | -51.5% |  | -37.65% |  | 107.8% | -74.71% | 17.78% | 1038% | 3.92% |  | -23.79% | 147% | 509.7% |  | -53.93% | 693.3%
| Z-Score (Absolute) |  | 0.67 | 0.45 | 6.74 | 0.47 | 6.74 | 0 | 36.71 | 7.93 | 0.5 | 0.4 | 0.67 | 0.3 | 6.74 | 0.75 | 0.57 | 0.1 | 4.97 | 0 | 0.15 | 0.2 | 1.03 | 3.65 | 0.67 | 0.42 | 4.97
| Analytical results for Ho in OREAS 234b (Certified Value 0.39 ppm) | [32 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 0.419 | 0.388 | 0.381 | 0.5 | 0.55 | 0.35 | [25 empty] | 1
| 1 | 2 | 0.417 | 0.386 | 0.38147 | 0.4 | 0.52 | 0.4 | [25 empty] | 20
| 1 | 3 | 0.408 | 0.435 | 0.36591 | 0.4 | 0.52 | 0.35 | [25 empty] | 16
| 1 | 4 | 0.409 | 0.408 | 0.36431 | 0.4 | 0.55 | 0.4 | [25 empty] | 0.39221
| 1 | 5 | 0.414 | 0.412 | 0.37337 | 0.4 | 0.5 | 0.3 | [25 empty] | 96
| 1 | 6 | 0.424 | 0.415 | 0.3652 | 0.4 | 0.51 | 0.3
| Mean |  | 0.41517 | 0.40733 | 0.37188 | 0.41667 | 0.525 | 0.35
| Median |  | 0.4155 | 0.41 | 0.36964 | 0.4 | 0.52 | 0.35
| Std Dev. |  | 0.0061128 | 0.018305 | 0.0079375 | 0.040825 | 0.020736 | 0.044721
| Rel.Std.Dev. |  | 1.472% | 4.494% | 2.134% | 9.798% | 3.95% | 12.78%
| PDM3 |  | 5.854% | 3.856% | -5.184% | 6.236% | 33.86% | -10.76%
| Z-Score (Absolute) |  | 0.12 | 0.12 | 1.19 | 0.16 | 3.43 | 1.84
| Analytical results for In in OREAS 234b (Certified Value 0.031 ppm) | [32 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.03 | 0.029149 | 0.03 | < 0.05 | 0.03 | < 5 | 0.03 | 0.034 | 0.033 | 0.028 | 0.029 | < 10 | < 0.05 | 0.58 | < 0.1 | 0.03086 | 0.029 | 0.03 | 0.03 | 0.03 | 0.023 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.03 | 0.028718 | 0.04 | < 0.05 | 0.03 | < 5 | 0.04 | 0.033 | 0.031 | 0.03 | 0.032 | < 10 | < 0.05 | 0.82 | < 0.1 | 0.02987 | 0.033 | 0.036 | 0.03 | 0.028 | 0.024 | [9 empty] | 21
| 1 | 3 | 0.03 | 0.03 | 0.027129 | 0.04 | < 0.05 | 0.03 | < 5 | 0.03 | 0.031 | 0.031 | 0.032 | 0.03 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.0348 | 0.033 | 0.038 | 0.03 | 0.028 | 0.024 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.03 | 0.028752 | 0.04 | < 0.05 | 0.03 | < 5 | 0.04 | 0.03 | 0.033 | 0.029 | 0.031 | < 10 | < 0.05 | 0.45 | < 0.1 | 0.03666 | 0.032 | 0.029 | 0.03 | 0.028 | 0.026 | [9 empty] | 0.030514
| 1 | 5 | 0.03 | 0.03 | 0.026532 | 0.04 | < 0.05 | 0.03 | < 5 | 0.04 | 0.033 | 0.031 | 0.029 | 0.03 | < 10 | < 0.05 | 1.67 | < 0.1 | 0.03447 | 0.032 | 0.028 | 0.03 | 0.029 | 0.027 | [9 empty] | 97
| 1 | 6 | 0.03 | 0.03 | 0.026665 | 0.03 | < 0.05 | 0.03 | < 5 | 0.03 | 0.03 | 0.03 | 0.031 | 0.031 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.03458 | 0.031 | 0.031 | 0.03 | 0.028 | 0.025
| Mean |  | 0.028333 | 0.03 | 0.027824 | 0.036667 |  | 0.03 |  | 0.035 | 0.031833 | 0.0315 | 0.029833 | 0.0305 |  |  | 0.88 |  | 0.03354 | 0.031667 | 0.032 | 0.03 | 0.0285 | 0.024833
| Median |  | 0.03 | 0.03 | 0.027924 | 0.04 |  | 0.03 |  | 0.035 | 0.032 | 0.031 | 0.0295 | 0.0305 |  |  | 0.7 |  | 0.034525 | 0.032 | 0.0305 | 0.03 | 0.028 | 0.0245
| Std Dev. |  | 0.0040825 | 0 | 0.0011758 | 0.005164 |  | 0 |  | 0.0054772 | 0.0017224 | 0.0012247 | 0.001472 | 0.0010488 |  |  | 0.54851 |  | 0.0026046 | 0.0015055 | 0.0040497 | 0 | 0.00083666 | 0.001472
| Rel.Std.Dev. |  | 14.41% | 0% | 4.226% | 14.08% |  | 0% |  | 15.65% | 5.411% | 3.888% | 4.934% | 3.439% |  |  | 62.33% |  | 7.766% | 4.754% | 12.66% | 0% | 2.936% | 5.927%
| PDM3 |  | -7.147% | -1.685% | -8.815% | 20.16% |  | -1.685% |  | 14.7% | 4.323% | 3.231% | -2.231% | -0.04624% |  |  | 2784% |  | 9.916% | 3.777% | 4.87% | -1.685% | -6.601% | -18.62%
| Z-Score (Absolute) |  | 0.69 | 0.26 | 0.82 | 1.47 | 1.55 | 0.26 | 639.51 | 1.04 | 0.22 | 0.13 | 0.3 | 0.13 | 1287.06 | 1.55 | 148.24 | 4.92 | 0.66 | 0.17 | 0.26 | 0.26 | 0.65 | 1.6
| Analytical results for Ir in OREAS 234b (Indicative Value < 0.003 ppm) | [32 empty] | Indicative
Ir | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.27 | 1.28 | [29 empty] | ppm
|  |  | AR*MS | AR*MS | [29 empty] | 3
|  |  | 25g | 0.5g | [29 empty] | 3
| 1 | 1 | N.A. | < 0.003 | [29 empty] | 1
| 1 | 2 | N.A. | < 0.003 | [29 empty] | 7
| 1 | 3 | N.A. | < 0.003 | [29 empty] | 16
| 1 | 4 | N.A. | < 0.003 | [29 empty] | < 0.003
| 1 | 5 | N.A. | < 0.003 | [29 empty] | 13
| 1 | 6 | N.A. | < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 234b (Certified Value 0.151 wt.%) | [32 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.15 | 0.14 | 0.19152 | 0.16128 | 0.15015 | 0.13 | 0.1635 | 0.165 | 0.1559 | 0.16 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15373 | 0.15 | 0.14 | 0.17269 | 0.153 | 0.13 | 0.15 | 0.14335 | 0.1529 | 0.1385 | 0.14 | 0.15 | 0.18 | 0.16 | 0.15 |  |  | 1
| 1 | 2 | 0.15 | 0.139 | 0.19342 | 0.16746 | 0.14888 | 0.13 | 0.161 | 0.173 | 0.1513 | 0.15 | 0.16 | 0.15 | 0.14 | 0.15 | 0.15191 | 0.15 | 0.14 | 0.16575 | 0.163 | 0.15 | 0.15 | 0.14769 | 0.14926 | 0.1365 | 0.14 | 0.15 | 0.17 | 0.16 | 0.15 |  |  | #N/A
| 1 | 3 | 0.16 | 0.139 | 0.19557 | 0.16077 | 0.15073 | 0.13 | 0.1602 | 0.175 | 0.1538 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15138 | 0.15 | 0.14 | 0.16559 | 0.156 | 0.15 | 0.15 | 0.14661 | 0.14795 | 0.138 | 0.14 | 0.15 | 0.17 | 0.16 | 0.15 |  |  | 16
| 1 | 4 | 0.16 | 0.139 | 0.19549 | 0.15917 | 0.14821 | 0.13 | 0.1594 | 0.174 | 0.1538 | 0.15 | 0.17 | 0.15 | 0.15 | 0.15 | 0.15195 | 0.15 | 0.13 | 0.16175 | 0.168 | 0.14 | 0.15 | 0.14582 | 0.14503 | 0.1365 | 0.14 | 0.15 | 0.17 | 0.15 | 0.15 |  |  | 0.15082
| 1 | 5 | 0.16 | 0.142 | 0.19259 | 0.16643 | 0.15003 | 0.13 | 0.1561 | 0.177 | 0.1601 | 0.15 | 0.16 | 0.15 | 0.15 | 0.15 | 0.15688 | 0.15 | 0.13 | 0.16612 | 0.154 | 0.14 | 0.15 | 0.14584 | 0.14896 | 0.1365 | 0.14 | 0.15 | 0.17 | 0.16 | 0.15 |  |  | 98
| 1 | 6 | 0.15 | 0.141 | 0.192 | 0.16294 | 0.15173 | 0.13 | 0.1619 | 0.176 | 0.1559 | 0.15 | 0.16 | 0.15 | 0.15 | 0.14 | 0.15828 | 0.15 | 0.14 | 0.17002 | 0.163 | 0.15 | 0.15 | 0.14446 | 0.14823 | 0.131 | 0.14 | 0.15 | 0.17 | 0.16 | 0.15
| Mean |  | 0.155 | 0.14 | 0.19343 | 0.16301 | 0.14996 | 0.13 | 0.16035 | 0.17333 | 0.15513 | 0.15167 | 0.15833 | 0.15 | 0.14833 | 0.14833 | 0.15402 | 0.15 | 0.13667 | 0.16699 | 0.1595 | 0.14333 | 0.15 | 0.14563 | 0.14872 | 0.13617 | 0.14 | 0.15 | 0.17167 | 0.15833 | 0.15
| Median |  | 0.155 | 0.1395 | 0.19301 | 0.16211 | 0.15009 | 0.13 | 0.1606 | 0.1745 | 0.15485 | 0.15 | 0.16 | 0.15 | 0.15 | 0.15 | 0.15284 | 0.15 | 0.14 | 0.16594 | 0.1595 | 0.145 | 0.15 | 0.14583 | 0.1486 | 0.1365 | 0.14 | 0.15 | 0.17 | 0.16 | 0.15
| Std Dev. |  | 0.0054772 | 0.0012649 | 0.0017437 | 0.0032942 | 0.0012645 | 0 | 0.0025193 | 0.0043205 | 0.0029696 | 0.0040825 | 0.0075277 | 0 | 0.0040825 | 0.0040825 | 0.0029021 | 0 | 0.005164 | 0.0038306 | 0.0060249 | 0.008165 | 0 | 0.0015413 | 0.0025405 | 0.0026771 | 0 | 0 | 0.0040825 | 0.0040825 | 0
| Rel.Std.Dev. |  | 3.534% | 0.9035% | 0.9015% | 2.021% | 0.8433% | 0% | 1.571% | 2.493% | 1.914% | 2.692% | 4.754% | 0% | 2.752% | 2.752% | 1.884% | 0% | 3.779% | 2.294% | 3.777% | 5.696% | 0% | 1.058% | 1.708% | 1.966% | 0% | 0% | 2.378% | 2.578% | 0%
| PDM3 |  | 2.771% | -7.175% | 28.25% | 8.08% | -0.5738% | -13.81% | 6.318% | 14.93% | 2.859% | 0.5604% | 4.981% | -0.5446% | -1.65% | -1.65% | 2.12% | -0.5446% | -9.385% | 10.72% | 5.754% | -4.965% | -0.5446% | -3.443% | -1.392% | -9.717% | -7.175% | -0.5446% | 13.82% | 4.981% | -0.5446%
| Z-Score (Absolute) |  | 0.51 | 1.01 | 4.39 | 1.32 | 0 | 2.02 | 1.05 | 2.36 | 0.52 | 0.17 | 0.84 | 0 | 0.17 | 0.17 | 0.41 | 0 | 1.35 | 1.72 | 0.96 | 0.67 | 0 | 0.44 | 0.13 | 1.4 | 1.01 | 0 | 2.19 | 0.84 | 0
| Analytical results for La in OREAS 234b (Certified Value 4.38 ppm) | [32 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 4.1 | 4.464 | 4.325 | 4.8 | 4.7831 | 4.4 | 1 | 4.75 | 4.395 | < 5 | 5 | 4.1 | 4.5 | 3.9 | 4.4 | 4 | 3.96 | 3.6 | 4.2 | 5 | 4.4 | 4.7 | 4.2 | 4.6 | 4.3 | [6 empty] | 1
| 1 | 2 | 4 | 4.383 | 3.732 | 4.87 | 4.8509 | 4.2 | 2 | 4.92 | 4.186 | < 5 | 5 | 4.1 | 4.5 | 4.3 | 4.3 | 4 | 3.9 | 4.1 | 4.3 | 4.5 | 4.4 | 4.6 | 4.4 | 4.5 | 4.4 | [6 empty] | 22
| 1 | 3 | 4.1 | 4.321 | 4.19 | 4.87 | 4.9092 | 4.4 | 1 | 4.96 | 4.391 | < 5 | 5 | 4.1 | 4.3 | 4.5 | 4.4 | 4 | 4.24 | 4 | 4.1 | 4.5 | 4.6 | 4.4 | 4.3 | 4.6 | 4.3 | [6 empty] | 16
| 1 | 4 | 4.2 | 4.427 | 3.9 | 4.71 | 4.8411 | 4.3 | < 1 | 4.94 | 4.311 | < 5 | 5 | 4.1 | 4.4 | 4.3 | 4.3 | 4 | 4.73 | 3.8 | 4.2 | 4.5 | 4.4 | 4.6 | 4.3 | 4.5 | 4.5 | [6 empty] | 4.3818
| 1 | 5 | 4.1 | 4.38 | 3.911 | 4.83 | 4.7423 | 4.3 | 1 | 4.88 | 4.499 | < 5 | 5 | 4.1 | 4.4 | 4.1 | 4.3 | 4 | 4.76 | 3.8 | 4.2 | 4.5 | 4.5 | 4.6 | 4.3 | 4.6 | 4.4 | [6 empty] | 99
| 1 | 6 | 4.1 | 4.375 | 4.036 | 4.89 | 4.9167 | 4.3 | 2 | 5.04 | 4.262 | < 5 | 5 | 4.1 | 4.3 | 4.2 | 4.2 | 4 | 4.63 | 3.9 | 4.3 | 4 | 4.4 | 4.7 | 4.4 | 4.5 | 4.3
| Mean |  | 4.1 | 4.3917 | 4.0157 | 4.8283 | 4.8405 | 4.3167 | 1.4 | 4.915 | 4.3407 |  | 5 | 4.1 | 4.4 | 4.2167 | 4.3167 | 4 | 4.37 | 3.8667 | 4.2167 | 4.5 | 4.45 | 4.6 | 4.3167 | 4.55 | 4.3667
| Median |  | 4.1 | 4.3815 | 3.9735 | 4.85 | 4.846 | 4.3 | 1 | 4.93 | 4.351 |  | 5 | 4.1 | 4.4 | 4.25 | 4.3 | 4 | 4.435 | 3.85 | 4.2 | 4.5 | 4.4 | 4.6 | 4.3 | 4.55 | 4.35
| Std Dev. |  | 0.063246 | 0.048915 | 0.21508 | 0.066458 | 0.068645 | 0.075277 | 0.54772 | 0.096695 | 0.11095 |  | 0 | 9.7295e-16 | 0.089443 | 0.20412 | 0.075277 | 0 | 0.38864 | 0.17512 | 0.075277 | 0.31623 | 0.083666 | 0.10954 | 0.075277 | 0.054772 | 0.08165
| Rel.Std.Dev. |  | 1.543% | 1.114% | 5.356% | 1.376% | 1.418% | 1.744% | 39.12% | 1.967% | 2.556% |  | 0% | 2.373e-14% | 2.033% | 4.841% | 1.744% | 0% | 8.893% | 4.529% | 1.785% | 7.027% | 1.88% | 2.381% | 1.744% | 1.204% | 1.87%
| PDM3 |  | -6.432% | 0.2244% | -8.357% | 10.19% | 10.47% | -1.487% | -68.05% | 12.17% | -0.9395% |  | 14.11% | -6.432% | 0.4145% | -3.769% | -1.487% | -8.714% | -0.2701% | -11.76% | -3.769% | 2.697% | 1.556% | 4.979% | -1.487% | 3.838% | -0.3462%
| Z-Score (Absolute) |  | 1.21 | 0.18 | 1.61 | 2.26 | 2.32 | 0.18 | NA | 2.67 | 0.06 | 8.82 | NA | 1.21 | 0.22 | 0.65 | 0.18 | NA | 0.08 | 2.32 | 0.65 | 0.7 | 0.46 | 1.17 | 0.18 | 0.93 | 0.06
Results from laboratories 1.08, 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 234b (Certified Value 10.3 ppm) | [32 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 1 | 9.96 | 16.29 | 3.8916 | 10.5179 | 10.3 | 6 | 12 | 10.3 | 12 | 13.5 | 11.2 | 10.8 | 11.1 | 11.1 | < 5 | 10.1635 | 8.9 | 10 | 9 | 11 | 9.8 | 11 | 11 | 10.1 | [6 empty] | 1
| 1 | 2 | 1 | 9.83 | 16.95 | 3.8209 | 10.4857 | 9.8 | 6 | 11.9 | 10.1 | 11 | 13.8 | 11.2 | 10.7 | 11.6 | 10.7 | < 5 | 8.7186 | 9 | 10 | 9 | 10.9 | 10.3 | 11 | 10.7 | 9.9 | [6 empty] | #N/A
| 1 | 3 | 1 | 10.04 | 16.94 | 3.8026 | 10.6169 | 10 | 6 | 11.8 | 10.3 | 11 | 13.3 | 11 | 10.6 | 11.4 | 10.7 | < 5 | 9.2232 | 9.8 | 10 | 10 | 11 | 9.6 | 11 | 10.6 | 9.8 | [6 empty] | 16
| 1 | 4 | 1 | 10.19 | 16.51 | 3.7829 | 10.3795 | 9.7 | 6 | 11.6 | 10.6 | 11 | 13.2 | 11 | 10.7 | 11.2 | 10.8 | < 5 | 8.1246 | 9.9 | 10 | 9 | 10.9 | 9.9 | 11 | 10.4 | 10.2 | [6 empty] | 10.3401
| 1 | 5 | < 1 | 10.13 | 15.9 | 3.858 | 10.3294 | 9.7 | 6 | 11.5 | 10.4 | 11 | 13.3 | 11 | 10.7 | 11.2 | 10.6 | < 5 | 10.3113 | 9.5 | 10 | 9 | 10.8 | 9.8 | 11 | 10.7 | 9.9 | [6 empty] | 100
| 1 | 6 | < 1 | 9.94 | 16.75 | 3.8986 | 10.4804 | 10 | 6 | 12 | 10.1 | 11 | 12.4 | 10.9 | 10.6 | 11.1 | 10.7 | < 5 | 8.71 | 10.2 | 10 | 9 | 10.9 | 9.6 | 11 | 10.5 | 9.6
| Mean |  | 1 | 10.015 | 16.5567 | 3.8424 | 10.4683 | 9.9167 | 6 | 11.8 | 10.3 | 11.1667 | 13.25 | 11.05 | 10.6833 | 11.2667 | 10.7667 |  | 9.2085 | 9.55 | 10 | 9.1667 | 10.9167 | 9.8333 | 11 | 10.65 | 9.9167
| Median |  | 1 | 10 | 16.63 | 3.8394 | 10.483 | 9.9 | 6 | 11.85 | 10.3 | 11 | 13.3 | 11 | 10.7 | 11.2 | 10.7 |  | 8.9709 | 9.65 | 10 | 9 | 10.9 | 9.8 | 11 | 10.65 | 9.9
| Std Dev. |  | 0 | 0.13217 | 0.41074 | 0.04776 | 0.10215 | 0.23166 | 0 | 0.20976 | 0.18974 | 0.40825 | 0.46797 | 0.12247 | 0.075277 | 0.19664 | 0.17512 |  | 0.87084 | 0.51672 | 0 | 0.40825 | 0.075277 | 0.2582 | 0 | 0.20736 | 0.2137
| Rel.Std.Dev. |  | 0% | 1.32% | 2.481% | 1.243% | 0.9758% | 2.336% | 0% | 1.778% | 1.842% | 3.656% | 3.532% | 1.108% | 0.7046% | 1.745% | 1.626% |  | 9.457% | 5.411% | 0% | 4.454% | 0.6896% | 2.626% | 0% | 1.947% | 2.155%
| PDM3 |  | -90.33% | -3.144% | 60.12% | -62.84% | 1.24% | -4.095% | -41.97% | 14.12% | -0.388% | 7.994% | 28.14% | 6.865% | 3.319% | 8.961% | 4.125% |  | -10.94% | -7.641% | -3.289% | -11.35% | 5.576% | -4.901% | 6.382% | 2.997% | -4.095%
| Z-Score (Absolute) |  | NA | 0.41 | 6.84 | 7.25 | 0.09 | 0.52 | NA | 1.57 | 0.09 | NA | 3.18 | 0.74 | 0.33 | 0.98 | 0.42 | 8.74 | 1.3 | 0.92 | NA | 1.35 | 0.59 | 0.61 | NA | 0.29 | 0.52
Results from laboratories 1.01, 1.08, 1.12, 1.24 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 234b (Certified Value 0.13 ppm) | [32 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | 1.32 | [24 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [24 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [24 empty] | 2
| 1 | 1 | 0.12 | 0.146 | 0.133 | 0.91374 | 0.2 | 0.1 | 0.13 | [24 empty] | 1
| 1 | 2 | 0.12 | 0.146 | 0.128 | 0.91033 | 0.1 | 0.1 | 0.13 | [24 empty] | 23
| 1 | 3 | 0.13 | 0.147 | 0.143 | 0.88537 | 0.1 | 0.1 | 0.12 | [24 empty] | 16
| 1 | 4 | 0.13 | 0.144 | 0.135 | 0.88456 | 0.1 | 0.1 | 0.13 | [24 empty] | 0.1337
| 1 | 5 | 0.12 | 0.149 | 0.136 | 0.88079 | 0.1 | 0.1 | 0.13 | [24 empty] | 101
| 1 | 6 | 0.13 | 0.146 | 0.138 | 0.87652 | 0.2 | < 0.1 | 0.13
| Mean |  | 0.125 | 0.14633 | 0.1355 | 0.89188 | 0.13333 | 0.1 | 0.12833
| Median |  | 0.125 | 0.146 | 0.1355 | 0.88497 | 0.1 | 0.1 | 0.13
| Std Dev. |  | 0.0054772 | 0.001633 | 0.00501 | 0.015959 | 0.05164 | 0 | 0.0040825
| Rel.Std.Dev. |  | 4.382% | 1.116% | 3.697% | 1.789% | 38.73% | 0% | 3.181%
| PDM3 |  | -6.507% | 9.449% | 1.346% | 567.1% | -0.2742% | -25.21% | -4.014%
| Z-Score (Absolute) |  | 0.67 | 1.05 | 0.18 | 61.38 | 0 | 3.37 | 0.4
| Analytical results for Mg in OREAS 234b (Certified Value 1.8 wt.%) | [32 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.87 | 1.749 | 1.7159 | 1.8946 | 1.8514 | 1.95 | 2.0022 | 1.98 | 1.8 | 1.8 | 1.73 | 1.81 | 1.78 | 1.75 | 1.7912 | 1.93 | 1.7 | 1.9398 | 1.9 | 1.58 | 1.79 | 2.2121 | 1.637 | 1.669 | 1.76 | 1.79 | 1.78 | 1.79 | 1.84 |  |  | 1
| 1 | 2 | 1.86 | 1.733 | 1.7336 | 1.9376 | 1.8419 | 1.85 | 1.9841 | 1.98 | 1.78 | 1.78 | 1.75 | 1.81 | 1.75 | 1.78 | 1.7717 | 1.86 | 1.69 | 1.8837 | 1.87 | 1.75 | 1.79 | 2.1135 | 1.6097 | 1.653 | 1.76 | 1.82 | 1.76 | 1.78 | 1.86 |  |  | #N/A
| 1 | 3 | 1.86 | 1.733 | 1.7744 | 1.9158 | 1.8665 | 1.88 | 1.978 | 1.99 | 1.81 | 1.8 | 1.74 | 1.81 | 1.74 | 1.78 | 1.7667 | 1.87 | 1.71 | 1.8942 | 1.88 | 1.78 | 1.8 | 2.3877 | 1.5167 | 1.683 | 1.77 | 1.76 | 1.76 | 1.79 | 1.86 |  |  | 16
| 1 | 4 | 1.9 | 1.744 | 1.7176 | 1.8951 | 1.8365 | 1.79 | 1.9479 | 1.96 | 1.82 | 1.8 | 1.73 | 1.8 | 1.76 | 1.77 | 1.7761 | 1.88 | 1.64 | 1.9384 | 1.92 | 1.7 | 1.81 | 2.1015 | 1.5667 | 1.677 | 1.76 | 1.8 | 1.77 | 1.78 | 1.86 |  |  | 1.8022
| 1 | 5 | 1.88 | 1.761 | 1.719 | 1.8955 | 1.8542 | 1.79 | 1.9419 | 1.96 | 1.87 | 1.87 | 1.73 | 1.81 | 1.76 | 1.74 | 1.8092 | 1.88 | 1.64 | 1.9564 | 1.87 | 1.66 | 1.79 | 2.184 | 1.6047 | 1.652 | 1.76 | 1.75 | 1.76 | 1.84 | 1.87 |  |  | 102
| 1 | 6 | 1.85 | 1.748 | 1.7317 | 1.9196 | 1.8452 | 1.82 | 2.0142 | 1.98 | 1.75 | 1.8 | 1.74 | 1.77 | 1.74 | 1.74 | 1.8446 | 1.86 | 1.74 | 1.9513 | 1.9 | 1.74 | 1.8 | 2.3468 | 1.606 | 1.642 | 1.76 | 1.76 | 1.77 | 1.79 | 1.84
| Mean |  | 1.87 | 1.7447 | 1.732 | 1.9097 | 1.8493 | 1.8467 | 1.978 | 1.975 | 1.805 | 1.8083 | 1.7367 | 1.8017 | 1.755 | 1.76 | 1.7932 | 1.88 | 1.6867 | 1.9273 | 1.89 | 1.7017 | 1.7967 | 2.2243 | 1.5901 | 1.6627 | 1.7617 | 1.78 | 1.7667 | 1.795 | 1.855
| Median |  | 1.865 | 1.746 | 1.7254 | 1.9057 | 1.8483 | 1.835 | 1.981 | 1.98 | 1.805 | 1.8 | 1.735 | 1.81 | 1.755 | 1.76 | 1.7836 | 1.875 | 1.695 | 1.9391 | 1.89 | 1.72 | 1.795 | 2.1981 | 1.6053 | 1.661 | 1.76 | 1.775 | 1.765 | 1.79 | 1.86
| Std Dev. |  | 0.017889 | 0.010671 | 0.022096 | 0.017643 | 0.010571 | 0.061536 | 0.028784 | 0.012247 | 0.040373 | 0.031252 | 0.008165 | 0.016021 | 0.015166 | 0.018974 | 0.029525 | 0.026077 | 0.039833 | 0.030653 | 0.02 | 0.072778 | 0.008165 | 0.11904 | 0.04239 | 0.016083 | 0.0040825 | 0.027568 | 0.008165 | 0.022583 | 0.012247
| Rel.Std.Dev. |  | 0.9566% | 0.6116% | 1.276% | 0.9239% | 0.5716% | 3.332% | 1.455% | 0.6201% | 2.237% | 1.728% | 0.4702% | 0.8892% | 0.8641% | 1.078% | 1.646% | 1.387% | 2.362% | 1.59% | 1.058% | 4.277% | 0.4545% | 5.352% | 2.666% | 0.9673% | 0.2317% | 1.549% | 0.4622% | 1.258% | 0.6602%
| PDM3 |  | 3.763% | -3.191% | -3.892% | 5.966% | 2.613% | 2.468% | 9.759% | 9.589% | 0.1564% | 0.3414% | -3.635% | -0.02856% | -2.618% | -2.341% | -0.4957% | 4.318% | -6.41% | 6.942% | 4.873% | -5.577% | -0.306% | 23.42% | -11.77% | -7.741% | -2.248% | -1.231% | -1.971% | -0.3985% | 2.931%
| Z-Score (Absolute) |  | 0.85 | 0.6 | 0.75 | 1.31 | 0.61 | 0.58 | 2.1 | 2.06 | 0.1 | 0.13 | 0.69 | 0.06 | 0.48 | 0.42 | 0.04 | 0.96 | 1.27 | 1.51 | 1.08 | 1.1 | 0 | 4.94 | 2.39 | 1.55 | 0.4 | 0.19 | 0.35 | 0.02 | 0.67
| Analytical results for Mn in OREAS 234b (Certified Value 0.082 wt.%) | [32 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.0872 | 0.08395 | 0.070565 | 0.087472 | 0.083565 | 0.0903 | 0.0922 | 0.0857 | 0.0834 | 0.0767 | 0.0934 | 0.0792 | 0.0796 | 0.078 | 0.079193 | 0.0848 | 0.0766 | 0.088602 | 0.0841 | 0.0707 | 0.0822 | 0.075801 | 0.085784 | 0.072 | 0.0762 | 0.0767 | 0.0847 | 0.0761 | 0.0828 |  |  | 1
| 1 | 2 | 0.0868 | 0.08302 | 0.069847 | 0.087506 | 0.082931 | 0.0862 | 0.0922 | 0.0867 | 0.082 | 0.0763 | 0.0964 | 0.079 | 0.0762 | 0.0804 | 0.07854 | 0.0827 | 0.0767 | 0.088351 | 0.0861 | 0.0782 | 0.0822 | 0.07708 | 0.085439 | 0.071 | 0.0766 | 0.0767 | 0.0871 | 0.0758 | 0.0837 |  |  | 24
| 1 | 3 | 0.0868 | 0.08328 | 0.071534 | 0.087581 | 0.083369 | 0.0889 | 0.0906 | 0.0869 | 0.0832 | 0.0755 | 0.0955 | 0.0786 | 0.0766 | 0.0801 | 0.078059 | 0.0839 | 0.0768 | 0.0891 | 0.0851 | 0.0788 | 0.083 | 0.076892 | 0.080889 | 0.073 | 0.0776 | 0.0757 | 0.0848 | 0.0763 | 0.0827 |  |  | 16
| 1 | 4 | 0.0888 | 0.08464 | 0.070179 | 0.086001 | 0.082442 | 0.0862 | 0.0898 | 0.0861 | 0.0843 | 0.0765 | 0.0958 | 0.0783 | 0.0772 | 0.0803 | 0.078602 | 0.085 | 0.0746 | 0.086155 | 0.0863 | 0.0753 | 0.0835 | 0.077316 | 0.081903 | 0.073 | 0.0767 | 0.078 | 0.0835 | 0.0758 | 0.0829 |  |  | 0.081662
| 1 | 5 | 0.0877 | 0.0851 | 0.070766 | 0.086664 | 0.08389 | 0.0848 | 0.0883 | 0.0858 | 0.0864 | 0.0758 | 0.0929 | 0.0785 | 0.078 | 0.0788 | 0.08028 | 0.0848 | 0.0741 | 0.091308 | 0.084 | 0.0741 | 0.0829 | 0.076484 | 0.084502 | 0.0715 | 0.0761 | 0.0748 | 0.0854 | 0.0783 | 0.0829 |  |  | 103
| 1 | 6 | 0.0859 | 0.08439 | 0.069268 | 0.087942 | 0.084527 | 0.0869 | 0.0914 | 0.0865 | 0.081 | 0.0764 | 0.093 | 0.0761 | 0.0768 | 0.0786 | 0.081387 | 0.0826 | 0.0788 | 0.08785 | 0.0865 | 0.0779 | 0.0831 | 0.076238 | 0.084768 | 0.07 | 0.0765 | 0.0782 | 0.0848 | 0.0762 | 0.0827
| Mean |  | 0.0872 | 0.084063 | 0.07036 | 0.087194 | 0.083454 | 0.087217 | 0.09075 | 0.086283 | 0.083383 | 0.0762 | 0.0945 | 0.078283 | 0.0774 | 0.079367 | 0.079343 | 0.083967 | 0.076267 | 0.088561 | 0.08535 | 0.075833 | 0.082817 | 0.076635 | 0.083881 | 0.07175 | 0.076617 | 0.076683 | 0.08505 | 0.076417 | 0.08295
| Median |  | 0.087 | 0.08417 | 0.070372 | 0.087489 | 0.083467 | 0.08655 | 0.091 | 0.0863 | 0.0833 | 0.07635 | 0.09445 | 0.07855 | 0.077 | 0.07945 | 0.078897 | 0.08435 | 0.07665 | 0.088477 | 0.0856 | 0.0766 | 0.08295 | 0.076688 | 0.084635 | 0.07175 | 0.07655 | 0.0767 | 0.0848 | 0.07615 | 0.08285
| Std Dev. |  | 0.00098184 | 0.00080361 | 0.00078403 | 0.00071945 | 0.00072897 | 0.0020173 | 0.0015202 | 0.0004916 | 0.0018745 | 0.00045607 | 0.0015697 | 0.0011197 | 0.0012394 | 0.001025 | 0.0012584 | 0.0010893 | 0.0017014 | 0.0016841 | 0.0011167 | 0.0031085 | 0.00051929 | 0.00056636 | 0.0020045 | 0.0011726 | 0.00053448 | 0.0013075 | 0.0011811 | 0.00094534 | 0.00037815
| Rel.Std.Dev. |  | 1.126% | 0.956% | 1.114% | 0.8251% | 0.8735% | 2.313% | 1.675% | 0.5697% | 2.248% | 0.5985% | 1.661% | 1.43% | 1.601% | 1.291% | 1.586% | 1.297% | 2.231% | 1.902% | 1.308% | 4.099% | 0.627% | 0.739% | 2.39% | 1.634% | 0.6976% | 1.705% | 1.389% | 1.237% | 0.4559%
| PDM3 |  | 6.781% | 2.94% | -13.84% | 6.774% | 2.194% | 6.802% | 11.13% | 5.659% | 2.108% | -6.689% | 15.72% | -4.138% | -5.219% | -2.811% | -2.839% | 2.822% | -6.607% | 8.448% | 4.516% | -7.138% | 1.414% | -6.156% | 2.717% | -12.14% | -6.179% | -6.097% | 4.149% | -6.423% | 1.577%
| Z-Score (Absolute) |  | 0.67 | 0.18 | 1.99 | 0.67 | 0.08 | 0.67 | 1.23 | 0.53 | 0.07 | 1.07 | 1.83 | 0.74 | 0.88 | 0.57 | 0.57 | 0.16 | 1.06 | 0.89 | 0.38 | 1.12 | 0.02 | 1 | 0.15 | 1.77 | 1 | 0.99 | 0.33 | 1.03 | 0
| Analytical results for Mo in OREAS 234b (Certified Value 2.41 ppm) | [32 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 3
| 1 | 1 | 2.39 | 2.4 | 2.4 | 1.9821 | 2.5151 | 2.67 | 2.22 | 2.3 | 2.4 | 2 | 2.5 | 2.44 | 2.37 | 2.27 | < 10 | 2.29 | 2 | < 5 | 2 | 1.96 | 3 | 2.44 | 2 | 2.3509 | 2.41 | 2.304 | 2.61 | 2.59 | 2.48 | 2.32 |  | 1
| 1 | 2 | 2.57 | 2.39 | 2.27 | 1.8657 | 2.5662 | 2.37 | 2.13 | 2.4 | 2.4 | 2 | 2.5 | 2.35 | 2.41 | 2.4 | < 10 | 2.23 | 2 | < 5 | 1.9 | 2.12 | 3 | 2.41 | 2 | 2.3897 | 2.39 | 2.275 | 2.57 | 2.6 | 2.41 | 2.41 |  | 25
| 1 | 3 | 2.56 | 2.41 | 2.51 | 1.8506 | 2.4318 | 2.36 | 2.22 | 2.4 | 2.6 | 2 | 2.4 | 2.29 | 2.45 | 2.54 | < 10 | 2.27 | 2 | < 5 | 2.1 | 2.17 | 3 | 3.11 | 2 | 2.2467 | 2.46 | 2.216 | 2.5 | 2.55 | 2.43 | 2.41 |  | 16
| 1 | 4 | 2.56 | 2.45 | 2.32 | 1.9055 | 2.4949 | 2.28 | 2.24 | 2.4 | 2.5 | 2 | 2.4 | 2.4 | 2.42 | 2.42 | < 10 | 2.23 | < 2 | < 5 | 1.8 | 2.16 | 3 | 2.95 | 2 | 2.4661 | 2.39 | 2.479 | 2.57 | 2.55 | 2.38 | 2.42 |  | 2.4136
| 1 | 5 | 2.53 | 2.41 | 2.65 | 1.9647 | 2.5446 | 2.43 | 2.21 | 2.4 | 2.4 | 2 | 2.4 | 2.36 | 2.39 | 2.4 | < 10 | 2.23 | < 2 | < 5 | 2.2 | 2.12 | 3 | 2.29 | 2 | 2.3979 | 2.48 | 2.421 | 2.5 | 2.59 | 2.51 | 2.43 |  | 104
| 1 | 6 | 2.39 | 2.4 | 2.45 | 1.8947 | 2.6176 | 2.34 | 2.2 | 2.4 | 2.5 | 2 | 2.4 | 2.42 | 2.39 | 2.39 | < 10 | 2.24 | 2 | < 5 | 1.8 | 2.22 | 3 | 2.69 | 2 | 2.5174 | 2.45 | 2.408 | 2.49 | 2.54 | 2.46 | 2.29
| Mean |  | 2.5 | 2.41 | 2.4333 | 1.9106 | 2.5284 | 2.4083 | 2.2033 | 2.3833 | 2.4667 | 2 | 2.4333 | 2.3767 | 2.405 | 2.4033 |  | 2.2483 | 2 |  | 1.9667 | 2.125 | 3 | 2.6483 | 2 | 2.3948 | 2.43 | 2.3505 | 2.54 | 2.57 | 2.445 | 2.38
| Median |  | 2.545 | 2.405 | 2.425 | 1.9001 | 2.5299 | 2.365 | 2.215 | 2.4 | 2.45 | 2 | 2.4 | 2.38 | 2.4 | 2.4 |  | 2.235 | 2 |  | 1.95 | 2.14 | 3 | 2.565 | 2 | 2.3938 | 2.43 | 2.356 | 2.535 | 2.57 | 2.445 | 2.41
| Std Dev. |  | 0.086255 | 0.020976 | 0.13692 | 0.052767 | 0.063707 | 0.13703 | 0.038297 | 0.040825 | 0.08165 | 0 | 0.05164 | 0.05465 | 0.028107 | 0.085946 |  | 0.025626 | 0 |  | 0.1633 | 0.088938 | 0 | 0.32695 | 0 | 0.093805 | 0.038471 | 0.10074 | 0.0498 | 0.026077 | 0.047645 | 0.05933
| Rel.Std.Dev. |  | 3.45% | 0.8704% | 5.627% | 2.762% | 2.52% | 5.69% | 1.738% | 1.713% | 3.31% | 0% | 2.122% | 2.299% | 1.169% | 3.576% |  | 1.14% | 0% |  | 8.303% | 4.185% | 0% | 12.35% | 0% | 3.917% | 1.583% | 4.286% | 1.961% | 1.015% | 1.949% | 2.493%
| PDM3 |  | 3.578% | -0.1505% | 0.8162% | -20.84% | 4.754% | -0.2195% | -8.713% | -1.255% | 2.197% | -17.14% | 0.8162% | -1.532% | -0.3576% | -0.4267% |  | -6.849% | -17.14% |  | -18.52% | -11.96% | 24.29% | 9.724% | -17.14% | -0.7815% | 0.6781% | -2.616% | 5.236% | 6.479% | 1.3% | -1.393%
| Z-Score (Absolute) |  | 1.05 | 0.01 | 0.28 | 5.79 | 1.38 | 0.01 | 2.39 | 0.3 | 0.67 | NA | 0.28 | 0.38 | 0.05 | 0.07 | 30.08 | 1.87 | NA | 1.05 | 5.14 | 3.3 | NA | 2.78 | NA | 0.17 | 0.24 | 0.68 | 1.52 | 1.87 | 0.42 | 0.34
Results from laboratories 1.12, 1.19, 1.24 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 234b (Certified Value 0.24 wt.%) | [32 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.25 | 0.237 | 1.0335 | 0.22088 | 0.23683 | 0.227 | 0.252 | 0.2 | 0.22 | 0.23 | 0.23 | 0.24 | 0.24 | 0.25 | 0.30064 | 0.24 | 0.23 | 0.23235 | 0.24 | 0.228 | 0.23 | 0.27 | 0.26362 | 0.233 | 0.23 | 0.22 | 0.28 | 0.25 | 0.27 |  |  | 1
| 1 | 2 | 0.25 | 0.236 | 1.1942 | 0.22713 | 0.2346 | 0.216 | 0.245 | 0.22 | 0.22 | 0.23 | 0.24 | 0.24 | 0.24 | 0.26 | 0.29752 | 0.24 | 0.22 | 0.23738 | 0.25 | 0.262 | 0.24 | 0.28 | 0.25699 | 0.228 | 0.23 | 0.23 | 0.28 | 0.25 | 0.27 |  |  | #N/A
| 1 | 3 | 0.25 | 0.235 | 1.1454 | 0.22959 | 0.23705 | 0.221 | 0.245 | 0.22 | 0.22 | 0.22 | 0.24 | 0.24 | 0.24 | 0.26 | 0.29805 | 0.24 | 0.23 | 0.23943 | 0.24 | 0.258 | 0.24 | 0.28 | 0.24803 | 0.23 | 0.23 | 0.22 | 0.28 | 0.25 | 0.27 |  |  | 16
| 1 | 4 | 0.25 | 0.236 | 1.1089 | 0.22109 | 0.23315 | 0.224 | 0.245 | 0.22 | 0.23 | 0.22 | 0.23 | 0.23 | 0.25 | 0.25 | 0.2995 | 0.24 | 0.22 | 0.23951 | 0.25 | 0.245 | 0.24 | 0.28 | 0.25286 | 0.23 | 0.23 | 0.22 | 0.28 | 0.25 | 0.27 |  |  | 0.2405
| 1 | 5 | 0.26 | 0.235 | 0.96597 | 0.23064 | 0.23707 | 0.215 | 0.237 | 0.22 | 0.23 | 0.23 | 0.23 | 0.23 | 0.24 | 0.25 | 0.30799 | 0.24 | 0.22 | 0.24839 | 0.24 | 0.238 | 0.23 | 0.28 | 0.25783 | 0.2265 | 0.22 | 0.22 | 0.28 | 0.26 | 0.27 |  |  | 105
| 1 | 6 | 0.25 | 0.237 | 1.1533 | 0.2261 | 0.23703 | 0.213 | 0.252 | 0.21 | 0.22 | 0.22 | 0.23 | 0.24 | 0.24 | 0.25 | 0.31252 | 0.24 | 0.23 | 0.2452 | 0.25 | 0.26 | 0.24 | 0.28 | 0.26125 | 0.2185 | 0.22 | 0.22 | 0.28 | 0.25 | 0.27
| Mean |  | 0.25167 | 0.236 | 1.1002 | 0.2259 | 0.23596 | 0.21933 | 0.246 | 0.215 | 0.22333 | 0.225 | 0.23333 | 0.23667 | 0.24167 | 0.25333 | 0.3027 | 0.24 | 0.225 | 0.24038 | 0.245 | 0.2485 | 0.23667 | 0.27833 | 0.25676 | 0.22767 | 0.22667 | 0.22167 | 0.28 | 0.25167 | 0.27
| Median |  | 0.25 | 0.236 | 1.1272 | 0.22662 | 0.23693 | 0.2185 | 0.245 | 0.22 | 0.22 | 0.225 | 0.23 | 0.24 | 0.24 | 0.25 | 0.30007 | 0.24 | 0.225 | 0.23947 | 0.245 | 0.2515 | 0.24 | 0.28 | 0.25741 | 0.229 | 0.23 | 0.22 | 0.28 | 0.25 | 0.27
| Std Dev. |  | 0.0040825 | 0.00089443 | 0.085142 | 0.0041466 | 0.0016767 | 0.0055377 | 0.0055857 | 0.0083666 | 0.005164 | 0.0054772 | 0.005164 | 0.005164 | 0.0040825 | 0.005164 | 0.0061199 | 0 | 0.0054772 | 0.0057019 | 0.0054772 | 0.013766 | 0.005164 | 0.0040825 | 0.0056538 | 0.0049967 | 0.005164 | 0.0040825 | 0 | 0.0040825 | 0
| Rel.Std.Dev. |  | 1.622% | 0.379% | 7.739% | 1.836% | 0.7106% | 2.525% | 2.271% | 3.891% | 2.312% | 2.434% | 2.213% | 2.182% | 1.689% | 2.038% | 2.022% | 0% | 2.434% | 2.372% | 2.236% | 5.54% | 2.182% | 1.467% | 2.202% | 2.195% | 2.278% | 1.842% | 0% | 1.622% | 0%
| PDM3 |  | 4.645% | -1.869% | 357.5% | -6.067% | -1.888% | -8.799% | 2.289% | -10.6% | -7.136% | -6.443% | -2.978% | -1.592% | 0.487% | 5.338% | 25.87% | -0.206% | -6.443% | -0.04997% | 1.873% | 3.328% | -1.592% | 15.73% | 6.764% | -5.334% | -5.75% | -7.829% | 16.43% | 4.645% | 12.27%
| Z-Score (Absolute) |  | 0.59 | 0.2 | 43.5 | 0.71 | 0.2 | 1.05 | 0.31 | 1.26 | 0.84 | 0.76 | 0.34 | 0.17 | 0.08 | 0.67 | 3.17 | 0 | 0.76 | 0.02 | 0.25 | 0.43 | 0.17 | 1.94 | 0.85 | 0.62 | 0.67 | 0.93 | 2.02 | 0.59 | 1.52
| Analytical results for Nb in OREAS 234b (Indicative Value 0.12 ppm) | [32 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | < 0.05 | 0.13 | 35.65 | 0.16887 | 0.1 | 0.35 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.08 | 0.15 | 0.09 | 0.06 | < 10 | < 0.2 | < 1 | 0.13 | 0.22 | 0.2 | 0.07 | 0.13 | [9 empty] | 1
| 1 | 2 | 0.19 | 0.14 | 34.5 | 0.13891 | 0.1 | 0.36 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.09 | 0.13 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.27 | 0.19 | 0.08 | 0.12 | [9 empty] | 8
| 1 | 3 | < 0.05 | 0.13 | 32.88 | 0.13994 | 0.1 | 0.35 | < 0.2 | 0.11 | < 5 | < 0.1 | 0.08 | 0.14 | 0.08 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.17 | 0.2 | 0.08 | 0.13 | [9 empty] | 16
| 1 | 4 | < 0.05 | 0.14 | 34.04 | 0.13814 | 0.1 | 0.36 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.09 | 0.13 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.12 | 0.2 | 0.09 | 0.13 | [9 empty] | 0.12197
| 1 | 5 | < 0.05 | 0.14 | 36.64 | 0.14915 | < 0.1 | 0.32 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.07 | 0.14 | 0.09 | 0.08 | < 10 | < 0.2 | < 1 | 0.1 | 0.12 | 0.2 | 0.08 | 0.13 | [9 empty] | 14
| 1 | 6 | < 0.05 | 0.14 | 32.14 | 0.14325 | 0.1 | 0.33 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.12 | 0.12 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.09 | 0.21 | 0.08 | 0.12
| Mean |  | 0.19 | 0.13667 | 34.3083 | 0.14638 | 0.1 | 0.345 |  | 0.095 |  |  | 0.088333 | 0.135 | 0.088333 | 0.07 |  |  |  | 0.11167 | 0.165 | 0.2 | 0.08 | 0.12667
| Median |  | 0.19 | 0.14 | 34.27 | 0.14159 | 0.1 | 0.35 |  | 0.09 |  |  | 0.085 | 0.135 | 0.09 | 0.07 |  |  |  | 0.11 | 0.145 | 0.2 | 0.08 | 0.13
| Std Dev. |  |  | 0.005164 | 1.6785 | 0.011735 | 0 | 0.016432 |  | 0.0083666 |  |  | 0.017224 | 0.010488 | 0.0040825 | 0.0063246 |  |  |  | 0.0098319 | 0.06892 | 0.0063246 | 0.0063246 | 0.005164
| Rel.Std.Dev. |  |  | 3.779% | 4.892% | 8.017% | 0% | 4.763% |  | 8.807% |  |  | 19.5% | 7.769% | 4.622% | 9.035% |  |  |  | 8.805% | 41.77% | 3.162% | 7.906% | 4.077%
| PDM3 |  | 55.78% | 12.05% | 28029% | 20.01% | -18.01% | 182.9% |  | -22.11% |  |  | -27.58% | 10.69% | -27.58% | -42.61% |  |  |  | -8.445% | 35.28% | 63.98% | -34.41% | 3.853%
| Z-Score (Absolute) |  | 1.28 | 0.34 | 658.69 | 0.52 | 0.53 | 4.35 | 0.37 | 0.47 | 45.87 | 1.33 | 0.59 | 0.31 | 0.59 | 0.95 | 94.03 | 0.37 | 7.34 | 0.14 | 0.88 | 1.56 | 0.75 | 0.14
| Analytical results for Nd in OREAS 234b (Certified Value 5.78 ppm) | [32 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 6.325 | 5.83 | 5.1274 | 6.17 | 6.3 | 6 | [25 empty] | 1
| 1 | 2 | 6.256 | 5.324 | 5.1012 | 6.12 | 6.08 | 5.5 | [25 empty] | 27
| 1 | 3 | 6.212 | 5.969 | 5.0712 | 6.06 | 6.09 | 5 | [25 empty] | 16
| 1 | 4 | 6.261 | 5.718 | 5.0498 | 6.01 | 6.09 | 5.5 | [25 empty] | 5.7785
| 1 | 5 | 6.268 | 5.611 | 5.1024 | 5.88 | 6.01 | 5.5 | [25 empty] | 106
| 1 | 6 | 6.333 | 5.799 | 5.0023 | 6.1 | 6 | 5.5
| Mean |  | 6.2758 | 5.7085 | 5.0757 | 6.0567 | 6.095 | 5.5
| Median |  | 6.2645 | 5.7585 | 5.0862 | 6.08 | 6.085 | 5.5
| Std Dev. |  | 0.045683 | 0.22282 | 0.04498 | 0.10211 | 0.10821 | 0.31623
| Rel.Std.Dev. |  | 0.7279% | 3.903% | 0.8862% | 1.686% | 1.775% | 5.75%
| PDM3 |  | 8.608% | -1.211% | -12.16% | 4.815% | 5.478% | -4.819%
| Z-Score (Absolute) |  | 0.89 | 0.39 | 1.83 | 0.39 | 0.48 | 0.87
| Analytical results for Ni in OREAS 234b (Certified Value 68 ppm) | [32 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 1
| 1 | 1 | 69.1 | 69 | 58.23 | 66.92 | 71.1288 | 68.4 | 69.3 | 75 | 68.2 | 62 | 61.5 | 71.2 | 68.1 | 67.2 | 67.023 | 69.9 | 59 | 72.0995 | 65 | 62.4 | 69.8 | 64.14 | 65.8168 | 64 | 70.9197 | 67.2 | 67 | 70.9 | 64.1 | 67 |  | 1
| 1 | 2 | 68.1 | 67 | 56.8 | 67.08 | 70.8207 | 66.8 | 67.4 | 75 | 65.8 | 61 | 63 | 69.7 | 67.8 | 71 | 70.223 | 68.1 | 59 | 70.8851 | 66 | 67 | 67.1 | 65.82 | 67.7764 | 63 | 73.2383 | 68.3 | 68.2 | 73.1 | 62.9 | 68.7 |  | 28
| 1 | 3 | 68.2 | 68 | 57.85 | 67.47 | 70.8971 | 69.2 | 71 | 73 | 69.1 | 60 | 61.3 | 69.7 | 62.9 | 71.8 | 65.444 | 69.1 | 59 | 73.2319 | 67 | 70.6 | 67.9 | 64.15 | 62.3116 | 64 | 71.718 | 68.9 | 66.7 | 71.5 | 64 | 67.1 |  | 16
| 1 | 4 | 70.2 | 68.7 | 57.31 | 65.87 | 69.983 | 65.4 | 70.2 | 74 | 70.7 | 61 | 60.8 | 70 | 67 | 71.1 | 68.436 | 69.1 | 58 | 73.0017 | 71 | 68.6 | 69.3 | 65.17 | 63.0108 | 64 | 68.7715 | 67.6 | 67 | 70.6 | 62.2 | 68.5 |  | 67.8733
| 1 | 5 | 68.6 | 66.4 | 56.75 | 66.44 | 71.6354 | 65.7 | 72.4 | 72 | 69.9 | 61 | 60.5 | 70 | 66.6 | 69.2 | 68.469 | 68.5 | 57 | 73.3356 | 73 | 65.1 | 70.5 | 64.02 | 68.7332 | 64 | 72.515 | 67.6 | 65.3 | 72.1 | 64.5 | 68.5 |  | 107
| 1 | 6 | 67.7 | 67.6 | 55.84 | 67.54 | 71.8108 | 67.4 | 69.7 | 76 | 68.2 | 60 | 60.2 | 68.2 | 63.2 | 68 | 70.919 | 67 | 61 | 73.3985 | 72 | 70.7 | 68.4 | 64.01 | 66.0192 | 65 | 71.3839 | 68.5 | 67.3 | 72.3 | 62.1 | 66.8
| Mean |  | 68.65 | 67.7833 | 57.13 | 66.8867 | 71.046 | 67.15 | 70 | 74.1667 | 68.65 | 60.8333 | 61.2167 | 69.8 | 65.9333 | 69.7167 | 68.419 | 68.6167 | 58.8333 | 72.6587 | 69 | 67.4 | 68.8333 | 64.5517 | 65.6113 | 64 | 71.4244 | 68.0167 | 66.9167 | 71.75 | 63.3 | 67.7667
| Median |  | 68.4 | 67.8 | 57.055 | 67 | 71.0129 | 67.1 | 69.95 | 74.5 | 68.65 | 61 | 61.05 | 69.85 | 66.8 | 70.1 | 68.4525 | 68.8 | 59 | 73.1168 | 69 | 67.8 | 68.85 | 64.145 | 65.918 | 64 | 71.551 | 67.95 | 67 | 71.8 | 63.45 | 67.8
| Std Dev. |  | 0.8961 | 0.99281 | 0.85769 | 0.63855 | 0.65454 | 1.491 | 1.6817 | 1.472 | 1.7026 | 0.75277 | 0.99883 | 0.96125 | 2.2993 | 1.8681 | 2.0153 | 1.0008 | 1.3292 | 0.99098 | 3.4059 | 3.2564 | 1.2644 | 0.7613 | 2.5414 | 0.63246 | 1.5401 | 0.64936 | 0.94534 | 0.93327 | 1.0373 | 0.88468
| Rel.Std.Dev. |  | 1.305% | 1.465% | 1.501% | 0.9547% | 0.9213% | 2.22% | 2.402% | 1.985% | 2.48% | 1.237% | 1.632% | 1.377% | 3.487% | 2.68% | 2.945% | 1.459% | 2.259% | 1.364% | 4.936% | 4.831% | 1.837% | 1.179% | 3.873% | 0.9882% | 2.156% | 0.9547% | 1.413% | 1.301% | 1.639% | 1.305%
| PDM3 |  | 1.144% | -0.1326% | -15.83% | -1.454% | 4.674% | -1.066% | 3.133% | 9.272% | 1.144% | -10.37% | -9.807% | 2.839% | -2.858% | 2.716% | 0.804% | 1.095% | -13.32% | 7.05% | 1.66% | -0.6973% | 1.414% | -4.894% | -3.333% | -5.707% | 5.232% | 0.2112% | -1.409% | 5.712% | -6.738% | -0.1571%
| Z-Score (Absolute) |  | 0.26 | 0.04 | 3.76 | 0.35 | 1.1 | 0.26 | 0.73 | 2.19 | 0.26 | 2.46 | 2.33 | 0.66 | 0.69 | 0.63 | 0.18 | 0.25 | 3.16 | 1.66 | 0.38 | 0.17 | 0.33 | 1.17 | 0.8 | 1.36 | 1.23 | 0.04 | 0.34 | 1.34 | 1.6 | 0.05
| Analytical results for P in OREAS 234b (Certified Value 0.042 wt.%) | [32 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.0441 | 0.0427 | 0.035134 | 0.043297 | 0.039 | 0.0419 | 0.0432 | 0.0424 | 0.0429 | 0.044 | 0.044 | 0.044 | 0.041 | 0.046032 | 0.045 | 0.0389 | 0.045769 | 0.043 | 0.037 | 0.04 | 0.043 | 0.042 | 0.042 | 0.0447 | 0.0408 | 0.042 | [5 empty] | 1
| 1 | 2 | 0.0446 | 0.0451 | 0.034694 | 0.043292 | 0.039 | 0.0415 | 0.0438 | 0.0411 | 0.042 | 0.046 | 0.044 | 0.043 | 0.042 | 0.045116 | 0.043 | 0.0384 | 0.045671 | 0.044 | 0.04 | 0.04 | 0.0425 | 0.042 | 0.044 | 0.0439 | 0.0407 | 0.043 | [5 empty] | #N/A
| 1 | 3 | 0.0441 | 0.0428 | 0.035421 | 0.043478 | 0.039 | 0.0415 | 0.0434 | 0.0381 | 0.0426 | 0.047 | 0.044 | 0.042 | 0.041 | 0.044606 | 0.044 | 0.0389 | 0.046454 | 0.044 | 0.04 | 0.04 | 0.042 | 0.041 | 0.042 | 0.0437 | 0.0411 | 0.042 | [5 empty] | 16
| 1 | 4 | 0.045 | 0.0448 | 0.035014 | 0.042447 | 0.039 | 0.041 | 0.0432 | 0.0412 | 0.0422 | 0.045 | 0.043 | 0.043 | 0.041 | 0.045346 | 0.043 | 0.0378 | 0.045623 | 0.045 | 0.041 | 0.04 | 0.042 | 0.041 | 0.043 | 0.0434 | 0.0407 | 0.041 | [5 empty] | 0.042484
| 1 | 5 | 0.0445 | 0.0454 | 0.035586 | 0.043059 | 0.04 | 0.041 | 0.0432 | 0.0415 | 0.0424 | 0.045 | 0.043 | 0.042 | 0.041 | 0.046245 | 0.043 | 0.0378 | 0.046655 | 0.044 | 0.04 | 0.04 | 0.042 | 0.041 | 0.042 | 0.0443 | 0.0419 | 0.042 | [5 empty] | 108
| 1 | 6 | 0.0441 | 0.0429 | 0.034292 | 0.043175 | 0.04 | 0.041 | 0.0433 | 0.0396 | 0.0423 | 0.043 | 0.042 | 0.042 | 0.04 | 0.047389 | 0.043 | 0.0401 | 0.046744 | 0.045 | 0.041 | 0.04 | 0.041 | 0.041 | 0.042 | 0.044 | 0.0408 | 0.042
| Mean |  | 0.0444 | 0.04395 | 0.035024 | 0.043125 | 0.039333 | 0.041317 | 0.04335 | 0.04065 | 0.0424 | 0.045 | 0.043333 | 0.042667 | 0.041 | 0.045789 | 0.0435 | 0.03865 | 0.046153 | 0.044167 | 0.039833 | 0.04 | 0.042083 | 0.041333 | 0.0425 | 0.044 | 0.041 | 0.042
| Median |  | 0.0443 | 0.04385 | 0.035074 | 0.043233 | 0.039 | 0.04125 | 0.04325 | 0.04115 | 0.04235 | 0.045 | 0.0435 | 0.0425 | 0.041 | 0.045689 | 0.043 | 0.03865 | 0.046112 | 0.044 | 0.04 | 0.04 | 0.042 | 0.041 | 0.042 | 0.04395 | 0.0408 | 0.042
| Std Dev. |  | 0.00036878 | 0.0012755 | 0.00047547 | 0.00036021 | 0.0005164 | 0.00037639 | 0.00023452 | 0.0015424 | 0.00031623 | 0.0014142 | 0.0008165 | 0.0008165 | 0.00063246 | 0.00098762 | 0.00083666 | 0.00086429 | 0.00052024 | 0.00075277 | 0.001472 | 0 | 0.00066458 | 0.0005164 | 0.00083666 | 0.00045607 | 0.00046476 | 0.00063246
| Rel.Std.Dev. |  | 0.8306% | 2.902% | 1.358% | 0.8353% | 1.313% | 0.911% | 0.541% | 3.794% | 0.7458% | 3.143% | 1.884% | 1.914% | 1.543% | 2.157% | 1.923% | 2.236% | 1.127% | 1.704% | 3.695% | 0% | 1.579% | 1.249% | 1.969% | 1.037% | 1.134% | 1.506%
| PDM3 |  | 4.511% | 3.451% | -17.56% | 1.508% | -7.416% | -2.747% | 2.039% | -4.316% | -0.1971% | 5.923% | 2% | 0.4306% | -3.492% | 7.78% | 2.392% | -9.024% | 8.636% | 3.961% | -6.239% | -5.846% | -0.9425% | -2.708% | 0.0383% | 3.569% | -3.492% | -1.139%
| Z-Score (Absolute) |  | 0.89 | 0.69 | 3.4 | 0.31 | 1.42 | 0.52 | 0.41 | 0.82 | 0.02 | 1.17 | 0.4 | 0.1 | 0.66 | 1.53 | 0.48 | 1.74 | 1.69 | 0.78 | 1.2 | 1.12 | 0.17 | 0.51 | 0.02 | 0.71 | 0.66 | 0.21
| Analytical results for Pb in OREAS 234b (Certified Value 25.9 ppm) | [32 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 2
| 1 | 1 | 27 | 25.6 | 24.82 | 25.61 | 25.1355 | 19.2 | 27 | 22.5 | 27.3 | 25 | 27.7 | 25.6 | 25.1 | 24.2 | 26.1 | 21 | 23.9246 | 24 | 21.91 | 25.4 | 27.8 | 26.8212 | 29 | 26.5471 | 25.3 | 33.164 | 25.7 | 25.4 | 26.2 | 26.1 |  | 1
| 1 | 2 | 26.9 | 25.6 | 24.97 | 25.92 | 25.8179 | 19.6 | 25.8 | 24.5 | 26.1 | 25 | 28.3 | 25.6 | 24.7 | 27 | 25.7 | 21 | 22.8406 | 23.3 | 24.28 | 26.1 | 29.12 | 26.8026 | 29 | 26.2877 | 25 | 32.289 | 26.1 | 26.3 | 25.8 | 27.2 |  | 29
| 1 | 3 | 27 | 25.1 | 25.31 | 26.31 | 25.757 | 19 | 26.1 | 24.4 | 26.4 | 24 | 27.8 | 25.5 | 23.9 | 27.1 | 25.9 | 22 | 25.1207 | 23.7 | 23.64 | 25.9 | 26.39 | 24.4218 | 28 | 26.2853 | 26.1 | 34.912 | 25.4 | 25.8 | 25.9 | 26.5 |  | 16
| 1 | 4 | 27.5 | 25.6 | 27.89 | 25.71 | 25.5319 | 18.3 | 26.2 | 24.3 | 27 | 25 | 27.9 | 25.8 | 24.7 | 27 | 25.3 | 21 | 25.0136 | 26.5 | 23.11 | 25.7 | 28.6 | 25.2309 | 29 | 26.7175 | 25.4 | 36.275 | 25.6 | 26 | 25.2 | 26.6 |  | 25.8888
| 1 | 5 | 27.1 | 25.8 | 27.17 | 25.94 | 25.8045 | 18.8 | 27.2 | 24.3 | 27.2 | 25 | 27.9 | 25.9 | 24.6 | 26.3 | 25.7 | 21 | 23.8615 | 26.4 | 23.16 | 25.8 | 25.64 | 25.6587 | 28 | 26.1061 | 25.4 | 31.72 | 25 | 26 | 26.2 | 26.8 |  | 109
| 1 | 6 | 27 | 25.6 | 25.84 | 26.18 | 26.9268 | 19.3 | 26.3 | 24.5 | 25.9 | 24 | 27.3 | 24.8 | 24.4 | 25.8 | 25.4 | 22 | 24.6041 | 24.8 | 24.22 | 26.1 | 28.85 | 25.6308 | 28 | 26.6703 | 25.5 | 33.731 | 25.8 | 25.7 | 25.9 | 26
| Mean |  | 27.0833 | 25.55 | 26 | 25.945 | 25.8289 | 19.0333 | 26.4333 | 24.0833 | 26.65 | 24.6667 | 27.8167 | 25.5333 | 24.5667 | 26.2333 | 25.6833 | 21.3333 | 24.2275 | 24.7833 | 23.3867 | 25.8333 | 27.7333 | 25.761 | 28.5 | 26.4357 | 25.45 | 33.6818 | 25.6 | 25.8667 | 25.8667 | 26.5333
| Median |  | 27 | 25.6 | 25.575 | 25.93 | 25.7808 | 19.1 | 26.25 | 24.35 | 26.7 | 25 | 27.85 | 25.6 | 24.65 | 26.65 | 25.7 | 21 | 24.2643 | 24.4 | 23.4 | 25.85 | 28.2 | 25.6448 | 28.5 | 26.4174 | 25.4 | 33.4475 | 25.65 | 25.9 | 25.9 | 26.55
| Std Dev. |  | 0.2137 | 0.23452 | 1.2567 | 0.26704 | 0.59672 | 0.45019 | 0.5465 | 0.78081 | 0.59582 | 0.5164 | 0.32506 | 0.38816 | 0.39833 | 1.1183 | 0.29944 | 0.5164 | 0.86101 | 1.3819 | 0.87922 | 0.26583 | 1.4221 | 0.92853 | 0.54772 | 0.24496 | 0.36194 | 1.691 | 0.37417 | 0.30768 | 0.36697 | 0.44572
| Rel.Std.Dev. |  | 0.789% | 0.9179% | 4.833% | 1.029% | 2.31% | 2.365% | 2.067% | 3.242% | 2.236% | 2.094% | 1.169% | 1.52% | 1.621% | 4.263% | 1.166% | 2.421% | 3.554% | 5.576% | 3.759% | 1.029% | 5.128% | 3.604% | 1.922% | 0.9266% | 1.422% | 5.02% | 1.462% | 1.189% | 1.419% | 1.68%
| PDM3 |  | 4.614% | -1.309% | 0.4296% | 0.2171% | -0.2312% | -26.48% | 2.103% | -6.974% | 2.94% | -4.721% | 7.447% | -1.373% | -5.107% | 1.331% | -0.7936% | -17.6% | -6.417% | -4.27% | -9.665% | -0.2142% | 7.125% | -0.4936% | 10.09% | 2.112% | -1.695% | 30.1% | -1.115% | -0.08544% | -0.08544% | 2.49%
| Z-Score (Absolute) |  | 1.29 | 0.29 | 0.17 | 0.12 | 0 | 7.02 | 0.62 | 1.8 | 0.85 | 1.2 | 2.05 | 0.31 | 1.3 | 0.42 | 0.15 | 4.64 | 1.66 | 1.08 | 2.52 | 0 | 1.96 | 0.07 | 2.75 | 0.62 | 0.39 | 8.1 | 0.24 | 0.04 | 0.04 | 0.72
| Analytical results for Pd in OREAS 234b (Indicative Value 16.1 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 1
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 1
| 1 | 1 | 19 | 12 | 14 | < 50 | < 20 | 31 | [25 empty] | 1
| 1 | 2 | 19 | 21 | 14 | < 50 | < 20 | 34 | [25 empty] | 1
| 1 | 3 | 13 | 14 | 13 | < 50 | 24.44 | 29 | [25 empty] | 16
| 1 | 4 | 19 | 11 | 17 | < 50 | < 20 | 41 | [25 empty] | 16.0556
| 1 | 5 | 14 | 23 | 17 | < 50 | < 20 | 28 | [25 empty] | 12
| 1 | 6 | 18 | 13 | 18 | < 50 | < 20 | 31
| Mean |  | 17 | 15.6667 | 15.5 |  | 24.44 | 32.3333
| Median |  | 18.5 | 13.5 | 15.5 |  | 24.44 | 31
| Std Dev. |  | 2.7568 | 5.0465 | 2.0736 |  |  | 4.7188
| Rel.Std.Dev. |  | 16.22% | 32.21% | 13.38% |  |  | 14.59%
| PDM3 |  | 5.882% | -2.422% | -3.46% |  | 52.22% | 101.4%
| Z-Score (Absolute) |  | 0.19 | 0.19 | 0.24 | 2.46 | 1.11 | 4.53
| Analytical results for Pr in OREAS 234b (Certified Value 1.23 ppm) | [32 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 1.311 | 1.219 | 1.1175 | 1.4 | 1.28 | 1.3 | [25 empty] | 1
| 1 | 2 | 1.31 | 1.091 | 1.1312 | 1.3 | 1.24 | 1.2 | [25 empty] | 30
| 1 | 3 | 1.279 | 1.217 | 1.1047 | 1.3 | 1.26 | 1.2 | [25 empty] | 16
| 1 | 4 | 1.308 | 1.156 | 1.1024 | 1.3 | 1.27 | 1.3 | [25 empty] | 1.2274
| 1 | 5 | 1.3 | 1.121 | 1.1188 | 1.3 | 1.24 | 1.2 | [25 empty] | 110
| 1 | 6 | 1.299 | 1.176 | 1.1132 | 1.3 | 1.22 | 1.1
| Mean |  | 1.3012 | 1.1633 | 1.1146 | 1.3167 | 1.2517 | 1.2167
| Median |  | 1.304 | 1.166 | 1.1154 | 1.3 | 1.25 | 1.2
| Std Dev. |  | 0.01199 | 0.051415 | 0.010511 | 0.040825 | 0.022286 | 0.075277
| Rel.Std.Dev. |  | 0.9215% | 4.42% | 0.943% | 3.101% | 1.781% | 6.187%
| PDM3 |  | 6.014% | -5.216% | -9.185% | 7.277% | 1.981% | -0.8708%
| Z-Score (Absolute) |  | 0.66 | 0.69 | 1.17 | 0.81 | 0.17 | 0.17
| Analytical results for Pt in OREAS 234b (Indicative Value 7.77 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 2
| 1 | 1 | 8 | 8 | 7 | < 10 | 9.82 | < 10 | [25 empty] | 1
| 1 | 2 | 9 | 7 | 6 | < 10 | 12.54 | < 10 | [25 empty] | 1
| 1 | 3 | 8 | 7 | 5 | < 10 | 8.74 | < 10 | [25 empty] | 16
| 1 | 4 | 8 | 7 | 7 | < 10 | 9.53 | < 10 | [25 empty] | 7.7663
| 1 | 5 | 8 | 8 | 5 | < 10 | 10.7 | < 10 | [25 empty] | 13
| 1 | 6 | 8 | 8 | 5 | < 10 | 6.06 | < 10
| Mean |  | 8.1667 | 7.5 | 5.8333 |  | 9.565
| Median |  | 8 | 7.5 | 5.5 |  | 9.675
| Std Dev. |  | 0.40825 | 0.54772 | 0.98319 |  | 2.1527
| Rel.Std.Dev. |  | 4.999% | 7.303% | 16.85% |  | 22.51%
| PDM3 |  | 5.156% | -3.428% | -24.89% |  | 23.16%
| Z-Score (Absolute) |  | 0.64 | 0.35 | 0.35 | 0.71 | 1.23 | 0.71
| Analytical results for Rb in OREAS 234b (Certified Value 6.33 ppm) | [32 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 5.9 | 6.38 | 5.9709 | 6.2129 | 5.1 | 7.24 | 6.85 | 6.34 | 6.8 | 5.9 | 6.6 | 5.5 | 6.1 | 6.24 | 6.3 | 6 | 6.4 | 6.1 | 6.2 | 6.9 | 5.8 | [10 empty] | 1
| 1 | 2 | 6.2 | 6.26 | 5.9218 | 6.286 | 4.9 | 7.18 | 7.17 | 6.29 | 6.8 | 5.8 | 6.3 | 5.7 | 5.9 | 7.23 | 6.2 | 6 | 6.6 | 5.9 | 6.4 | 6.7 | 6 | [10 empty] | 31
| 1 | 3 | 6 | 6.27 | 5.8573 | 6.2753 | 5.1 | 7.29 | 7.22 | 6.51 | 6.9 | 5.9 | 6.4 | 6 | 6.1 | 6.21 | 6.5 | 6 | 6.6 | 5.9 | 6.3 | 6.8 | 5.8 | [10 empty] | 16
| 1 | 4 | 6 | 6.44 | 5.7907 | 6.2354 | 4.9 | 7.37 | 7.17 | 6.37 | 7.1 | 5.9 | 6.3 | 5.8 | 6.1 | 6.86 | 6.6 | 6 | 6.5 | 6.1 | 6.1 | 6.6 | 6 | [10 empty] | 6.3303
| 1 | 5 | 6.2 | 6.39 | 5.8634 | 6.3114 | 5 | 7.18 | 7.18 | 6.88 | 6.7 | 5.8 | 6.3 | 5.7 | 6 | 6.88 | 6.6 | 6 | 6.3 | 5.9 | 6.2 | 6.8 | 5.9 | [10 empty] | 111
| 1 | 6 | 5.9 | 6.28 | 5.9956 | 6.3276 | 5 | 7.23 | 7.2 | 6.36 | 6.7 | 5.8 | 6.1 | 5.6 | 6 | 6.68 | 6.5 | 5.5 | 6.6 | 6.2 | 6.4 | 6.5 | 5.6
| Mean |  | 6.0333 | 6.3367 | 5.8999 | 6.2748 | 5 | 7.2483 | 7.1317 | 6.4583 | 6.8333 | 5.85 | 6.3333 | 5.7167 | 6.0333 | 6.6833 | 6.45 | 5.9167 | 6.5 | 6.0167 | 6.2667 | 6.7167 | 5.85
| Median |  | 6 | 6.33 | 5.8926 | 6.2806 | 5 | 7.235 | 7.175 | 6.365 | 6.8 | 5.85 | 6.3 | 5.7 | 6.05 | 6.77 | 6.5 | 6 | 6.55 | 6 | 6.25 | 6.75 | 5.85
| Std Dev. |  | 0.13663 | 0.07607 | 0.07715 | 0.043912 | 0.089443 | 0.072503 | 0.13934 | 0.21922 | 0.15055 | 0.054772 | 0.1633 | 0.17224 | 0.08165 | 0.39732 | 0.16432 | 0.20412 | 0.12649 | 0.13292 | 0.12111 | 0.1472 | 0.15166
| Rel.Std.Dev. |  | 2.265% | 1.2% | 1.308% | 0.6998% | 1.789% | 1% | 1.954% | 3.394% | 2.203% | 0.9363% | 2.578% | 3.013% | 1.353% | 5.945% | 2.548% | 3.45% | 1.946% | 2.209% | 1.933% | 2.192% | 2.592%
| PDM3 |  | -4.69% | 0.1013% | -6.798% | -0.8765% | -21.01% | 14.5% | 12.66% | 2.023% | 7.947% | -7.587% | 0.04868% | -9.693% | -4.69% | 5.578% | 1.892% | -6.533% | 2.682% | -4.954% | -1.004% | 6.104% | -7.587%
| Z-Score (Absolute) |  | 0.45 | 0.12 | 0.71 | 0 | 2.4 | 1.83 | 1.61 | 0.35 | 1.05 | 0.8 | 0.11 | 1.05 | 0.45 | 0.77 | 0.33 | 0.67 | 0.42 | 0.49 | 0.02 | 0.83 | 0.8
| Analytical results for Re in OREAS 234b (Indicative Value 0.003 ppm) | [32 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.008 | < 0.002 | 0.002 | 0.004 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.007 | < 0.002 | 0.002 | 0.004 | [13 empty] | 3
| 1 | 3 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.009 | < 0.002 | 0.003 | 0.003 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.001 | < 0.005 | 0.002 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.007 | < 0.002 | 0.003 | 0.002 | [13 empty] | 0.0025606
| 1 | 5 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.007 | < 0.002 | 0.002 | 0.004 | [13 empty] | 14
| 1 | 6 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.001 | < 0.005 | 0.001 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.009 | < 0.002 | 0.002 | 0.003
| Mean |  |  | 0.002 | 0.002 |  | 0.0018333 |  | 0.002 | 0.002 | 0.0016667 | 0.002 |  |  |  | 0.002 | 0.0078333 |  | 0.0023333 | 0.0033333
| Median |  |  | 0.002 | 0.002 |  | 0.002 |  | 0.002 | 0.002 | 0.002 | 0.002 |  |  |  | 0.002 | 0.0075 |  | 0.002 | 0.0035
| Std Dev. |  |  | 0 | 0 |  | 0.00075277 |  | 0.00063246 | 0 | 0.0005164 | 0 |  |  |  | 0 | 0.00098319 |  | 0.0005164 | 0.0008165
| Rel.Std.Dev. |  |  | 0% | 0% |  | 41.06% |  | 31.62% | 0% | 30.98% | 0% |  |  |  | 0% | 12.55% |  | 22.13% | 24.49%
| PDM3 |  |  | -21.89% | -21.89% |  | -28.4% |  | -21.89% | -21.89% | -34.91% | -21.89% |  |  |  | -21.89% | 205.9% |  | -8.876% | 30.18%
| Z-Score (Absolute) |  | 77.41 | 0.27 | 0.27 | 36.95 | 0.54 | 0.54 | 0.27 | 0.27 | 0.81 | 0.27 | 4042.35 | 36.95 | 36.95 | 0.27 | 9.17 | 1.89 | 0.27 | 1.89
| Analytical results for S in OREAS 234b (Certified Value 0.402 wt.%) | [32 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.42 | 0.39 | 0.44393 | 0.40498 | 0.43144 | 0.38 | 0.402 | 0.38 | 0.4 | 0.41 | 0.4 | 0.4 | 0.4 | 0.4 | 0.43 | 0.35 | 0.37793 | 0.39 | 0.31 | 0.41 | 0.4108 | 0.41 | 0.38 | 0.39 | 0.43 | 0.4 | 0.41 |  |  |  |  | 1
| 1 | 2 | 0.42 | 0.39 | 0.42666 | 0.40497 | 0.42435 | 0.376 | 0.402 | 0.41 | 0.38 | 0.39 | 0.41 | 0.41 | 0.4 | 0.41 | 0.41 | 0.35 | 0.37028 | 0.4 | 0.34 | 0.4 | 0.40776 | 0.4035 | 0.37 | 0.41 | 0.42 | 0.39 | 0.42 |  |  |  |  | 17
| 1 | 3 | 0.41 | 0.38 | 0.43619 | 0.40164 | 0.43202 | 0.374 | 0.394 | 0.41 | 0.39 | 0.4 | 0.4 | 0.41 | 0.37 | 0.41 | 0.42 | 0.35 | 0.37825 | 0.4 | 0.35 | 0.4 | 0.38594 | 0.405 | 0.38 | 0.38 | 0.42 | 0.39 | 0.41 |  |  |  |  | 16
| 1 | 4 | 0.42 | 0.39 | 0.41177 | 0.40297 | 0.41962 | 0.378 | 0.4 | 0.4 | 0.39 | 0.39 | 0.4 | 0.4 | 0.4 | 0.4 | 0.42 | 0.34 | 0.38151 | 0.41 | 0.35 | 0.4 | 0.39324 | 0.4065 | 0.38 | 0.41 | 0.42 | 0.39 | 0.41 |  |  |  |  | 0.40163
| 1 | 5 | 0.42 | 0.39 | 0.42784 | 0.40333 | 0.42418 | 0.385 | 0.396 | 0.41 | 0.41 | 0.41 | 0.4 | 0.4 | 0.38 | 0.4 | 0.42 | 0.34 | 0.3964 | 0.4 | 0.33 | 0.4 | 0.40592 | 0.4045 | 0.37 | 0.37 | 0.42 | 0.4 | 0.41 |  |  |  |  | 112
| 1 | 6 | 0.41 | 0.38 | 0.47067 | 0.40677 | 0.41896 | 0.393 | 0.402 | 0.4 | 0.38 | 0.39 | 0.4 | 0.41 | 0.39 | 0.4 | 0.41 | 0.36 | 0.39242 | 0.41 | 0.33 | 0.4 | 0.40202 | 0.399 | 0.37 | 0.4 | 0.43 | 0.39 | 0.41
| Mean |  | 0.41667 | 0.38667 | 0.43618 | 0.40411 | 0.42509 | 0.381 | 0.39933 | 0.40167 | 0.39167 | 0.39833 | 0.40167 | 0.405 | 0.39 | 0.40333 | 0.41833 | 0.34833 | 0.3828 | 0.40167 | 0.335 | 0.40167 | 0.40095 | 0.40475 | 0.375 | 0.39333 | 0.42333 | 0.39333 | 0.41167
| Median |  | 0.42 | 0.39 | 0.43201 | 0.40415 | 0.42426 | 0.379 | 0.401 | 0.405 | 0.39 | 0.395 | 0.4 | 0.405 | 0.395 | 0.4 | 0.42 | 0.35 | 0.37988 | 0.4 | 0.335 | 0.4 | 0.40397 | 0.40475 | 0.375 | 0.395 | 0.42 | 0.39 | 0.41
| Std Dev. |  | 0.005164 | 0.005164 | 0.020028 | 0.0018198 | 0.005609 | 0.0069857 | 0.0035024 | 0.01169 | 0.01169 | 0.0098319 | 0.0040825 | 0.0054772 | 0.012649 | 0.005164 | 0.0075277 | 0.0075277 | 0.0097998 | 0.0075277 | 0.015166 | 0.0040825 | 0.0095297 | 0.0036159 | 0.0054772 | 0.01633 | 0.005164 | 0.005164 | 0.0040825
| Rel.Std.Dev. |  | 1.239% | 1.336% | 4.592% | 0.4503% | 1.319% | 1.834% | 0.8771% | 2.91% | 2.985% | 2.468% | 1.016% | 1.352% | 3.243% | 1.28% | 1.799% | 2.161% | 2.56% | 1.874% | 4.527% | 1.016% | 2.377% | 0.8934% | 1.461% | 4.152% | 1.22% | 1.313% | 0.9917%
| PDM3 |  | 3.745% | -3.725% | 8.603% | 0.6186% | 5.843% | -5.136% | -0.5708% | 0.01017% | -2.48% | -0.8198% | 0.01017% | 0.8401% | -2.895% | 0.4251% | 4.16% | -13.27% | -4.687% | 0.01017% | -16.59% | 0.01017% | -0.1691% | 0.7779% | -6.63% | -2.065% | 5.405% | -2.065% | 2.5%
| Z-Score (Absolute) |  | 1.01 | 1.01 | 2.33 | 0.16 | 1.58 | 1.39 | 0.16 | 0 | 0.67 | 0.22 | 0 | 0.22 | 0.79 | 0.11 | 1.12 | 3.6 | 1.27 | 0 | 4.5 | 0 | 0.05 | 0.21 | 1.8 | 0.56 | 1.46 | 0.56 | 0.67
| Analytical results for Sb in OREAS 234b (Certified Value 0.87 ppm) | [32 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 1.18 | 1.01 | 4.07 | 1.0004 | 1.19 | 1.11 | 0.56 | 0.47 | < 5 | 1.06 | 0.85 | 1.03 | 0.98 | 80.903 | 0.86 | < 5 | < 5 | 0.93 | 0.65 | 0.54 | 5.91 | 1 | 1.1624 | 1.04 | 0.22 | 0.61 | 0.51 | 0.59 | 0.7 |  |  | 1
| 1 | 2 | 1.6 | 0.98 | 5.15 | 1.0172 | 0.87 | 1.07 | 0.28 | 0.45 | < 5 | 1.07 | 0.87 | 1.09 | 1.03 | 82.561 | 0.9 | < 5 | < 5 | 1.03 | 0.71 | 0.55 | 7.04 | 0.5 | 1.0858 | 1.02 | 0.222 | 0.62 | 0.53 | 0.58 | 0.75 |  |  | 33
| 1 | 3 | 1.19 | 1.01 | 4.43 | 1.0005 | 0.93 | 1.1 | 0.31 | 0.46 | < 5 | 1.09 | 0.93 | 1.08 | 1.06 | 81.432 | 0.9 | < 5 | < 5 | 1.04 | 0.65 | 0.58 | 6.8 | 0.5 | 1.1736 | 1.04 | 0.245 | 0.64 | 0.52 | 0.59 | 0.72 |  |  | 16
| 1 | 4 | 1.2 | 1.01 | 5.03 | 0.97088 | 0.84 | 1.08 | 0.28 | 0.47 | < 5 | 1.14 | 0.92 | 1.13 | 1.01 | 81.221 | 0.9 | < 5 | < 5 | 0.9 | 0.71 | 0.61 | 7.55 | 1 | 1.1522 | 1.02 | 0.229 | 0.62 | 0.5 | 0.56 | 0.71 |  |  | 0.8721
| 1 | 5 | 1.2 | 1 | 4.56 | 1.0033 | 0.85 | 1.07 | 0.26 | 0.49 | < 5 | 1.04 | 0.89 | 1.08 | 1 | 81.075 | 0.89 | < 5 | < 5 | 1.07 | 0.65 | 0.61 | 6.95 | 1 | 1.1438 | 1.04 | 0.223 | 0.6 | 0.51 | 0.6 | 0.73 |  |  | 113
| 1 | 6 | 1.19 | 1.01 | 3.79 | 1.0404 | 0.85 | 1.08 | 0.27 | 0.45 | < 5 | 1.02 | 0.91 | 1.12 | 0.98 | 85.371 | 0.86 | < 5 | < 5 | 1 | 0.76 | 0.62 | 8.05 | 1 | 1.2316 | 0.99 | 0.228 | 0.59 | 0.52 | 0.6 | 0.64
| Mean |  | 1.26 | 1.0033 | 4.505 | 1.0054 | 0.92167 | 1.085 | 0.32667 | 0.465 |  | 1.07 | 0.895 | 1.0883 | 1.01 | 82.0938 | 0.885 |  |  | 0.995 | 0.68833 | 0.585 | 7.05 | 0.83333 | 1.1582 | 1.025 | 0.22783 | 0.61333 | 0.515 | 0.58667 | 0.70833
| Median |  | 1.195 | 1.01 | 4.495 | 1.0019 | 0.86 | 1.08 | 0.28 | 0.465 |  | 1.065 | 0.9 | 1.085 | 1.005 | 81.3265 | 0.895 |  |  | 1.015 | 0.68 | 0.595 | 6.995 | 1 | 1.1573 | 1.03 | 0.2255 | 0.615 | 0.515 | 0.59 | 0.715
| Std Dev. |  | 0.16673 | 0.012111 | 0.52929 | 0.022852 | 0.13541 | 0.016432 | 0.11553 | 0.015166 |  | 0.041952 | 0.030822 | 0.035449 | 0.030984 | 1.7096 | 0.019748 |  |  | 0.066558 | 0.045789 | 0.033912 | 0.72418 | 0.2582 | 0.047166 | 0.019748 | 0.0091086 | 0.017512 | 0.010488 | 0.015055 | 0.037639
| Rel.Std.Dev. |  | 13.23% | 1.207% | 11.75% | 2.273% | 14.69% | 1.514% | 35.37% | 3.261% |  | 3.921% | 3.444% | 3.257% | 3.068% | 2.083% | 2.231% |  |  | 6.689% | 6.652% | 5.797% | 10.27% | 30.98% | 4.072% | 1.927% | 3.998% | 2.855% | 2.037% | 2.566% | 5.314%
| PDM3 |  | 44.48% | 15.05% | 416.6% | 15.29% | 5.683% | 24.41% | -62.54% | -46.68% |  | 22.69% | 2.626% | 24.79% | 15.81% | 9313% | 1.479% |  |  | 14.09% | -21.07% | -32.92% | 708.4% | -4.445% | 32.81% | 17.53% | -73.88% | -29.67% | -40.95% | -32.73% | -18.78%
| Z-Score (Absolute) |  | 0.56 | 0 | 7.71 | 0 | 0.18 | 0.18 | 1.49 | 1.19 | 3.3 | 0.14 | 0.24 | 0.18 | 0.01 | 178.66 | 0.26 | 3.3 | 3.3 | 0.02 | 0.7 | 0.92 | 13.32 | NA | 0.34 | 0.05 | 1.71 | 0.86 | 1.08 | 0.92 | 0.65
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 234b (Certified Value 4.96 ppm) | [32 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 4.2 | 4.93 | 4.659 | 6.26 | 5.0806 | 6.5 | 7.2 | 6 | 5.8 | 5 | 4.5 | 5.1 | 5.6 | 4.4 | 4.7 | 4 | 6 | 3.6 | 4.6 | 4 | 5 | 1.5 | 4.6 | 4.8 | 3.6 | [6 empty] | 1
| 1 | 2 | 4.1 | 4.98 | 4.88 | 6.32 | 4.9802 | 6.2 | 7.3 | 6 | 5.6 | 5 | 4.6 | 5 | 5.4 | 4.9 | 4.7 | 4 | 5 | 4.2 | 4.6 | 4 | 5 | 1.6 | 4.9 | 4.9 | 3.7 | [6 empty] | 34
| 1 | 3 | 4.3 | 4.95 | 5.796 | 6.46 | 5.0251 | 6.2 | 6.9 | 6 | 5.8 | 5 | 4.5 | 5 | 5.5 | 5.1 | 4.8 | 4 | 5 | 4.2 | 4.7 | 4 | 5.1 | 1.6 | 5.2 | 4.9 | 3.6 | [6 empty] | 16
| 1 | 4 | 4.4 | 4.94 | 5.217 | 6.13 | 5.0063 | 6.2 | 7 | 6 | 5.7 | 5 | 4.6 | 5 | 5.4 | 4.9 | 4.8 | 4 | 6 | 3.7 | 4.7 | 4 | 5 | 1.5 | 4.9 | 4.8 | 3.7 | [6 empty] | 4.9578
| 1 | 5 | 4.5 | 5 | 5.51 | 6.32 | 5.0802 | 5.9 | 6.2 | 6 | 6 | 5 | 4.4 | 4.9 | 5.3 | 4.7 | 4.9 | 4 | 6 | 3.8 | 4.7 | 4 | 5 | 1.4 | 4.6 | 5 | 3.8 | [6 empty] | 114
| 1 | 6 | 4.3 | 4.93 | 5.283 | 6.25 | 5.0507 | 6.1 | 6.5 | 6 | 5.8 | 5 | 4.3 | 5 | 5.2 | 4.8 | 4.7 | 4 | 6 | 4.2 | 4.7 | 3 | 4.9 | 1.3 | 4.9 | 4.9 | 3.6
| Mean |  | 4.3 | 4.955 | 5.2242 | 6.29 | 5.0372 | 6.1833 | 6.85 | 6 | 5.7833 | 5 | 4.4833 | 5 | 5.4 | 4.8 | 4.7667 | 4 | 5.6667 | 3.95 | 4.6667 | 3.8333 | 5 | 1.4833 | 4.85 | 4.8833 | 3.6667
| Median |  | 4.3 | 4.945 | 5.25 | 6.29 | 5.0379 | 6.2 | 6.95 | 6 | 5.8 | 5 | 4.5 | 5 | 5.4 | 4.85 | 4.75 | 4 | 6 | 4 | 4.7 | 4 | 5 | 1.5 | 4.9 | 4.9 | 3.65
| Std Dev. |  | 0.14142 | 0.02881 | 0.41233 | 0.10844 | 0.040653 | 0.19408 | 0.42308 | 0 | 0.13292 | 0 | 0.1169 | 0.063246 | 0.14142 | 0.23664 | 0.08165 | 0 | 0.5164 | 0.28107 | 0.05164 | 0.40825 | 0.063246 | 0.1169 | 0.22583 | 0.075277 | 0.08165
| Rel.Std.Dev. |  | 3.289% | 0.5814% | 7.893% | 1.724% | 0.8071% | 3.139% | 6.176% | 0% | 2.298% | 0% | 2.608% | 1.265% | 2.619% | 4.93% | 1.713% | 0% | 9.113% | 7.116% | 1.107% | 10.65% | 1.265% | 7.881% | 4.656% | 1.542% | 2.227%
| PDM3 |  | -13.27% | -0.05565% | 5.374% | 26.87% | 1.602% | 24.72% | 38.17% | 21.02% | 16.65% | 0.852% | -9.569% | 0.852% | 8.92% | -3.182% | -3.854% | -19.32% | 14.3% | -20.33% | -5.871% | -22.68% | 0.852% | -70.08% | -2.174% | -1.501% | -26.04%
| Z-Score (Absolute) |  | 1.13 | 0.07 | 0.56 | 2.5 | 0.21 | 2.3 | 3.51 | NA | 1.57 | NA | 0.79 | 0.15 | 0.88 | 0.22 | 0.28 | NA | NA | 1.76 | 0.46 | NA | 0.15 | 6.25 | 0.13 | 0.07 | 2.28
| Analytical results for Se in OREAS 234b (Indicative Value 0.66 ppm) | [32 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | 0.49 | < 1 | < 1.5 | 0.6 | 1 | 1 | < 1 | < 5 | 2 | 0.8 | 0.6 | 0.8 | 0.8 | < 5 | < 5 | < 1 | 0.5 | 3.08 | < 1 | 0.46473 | 0.8 | 0.475 | 2.6 | < 1 | 0.5 | 0.5 |  |  |  |  | 1
| 1 | 2 | < 1 | 0.45 | < 1 | < 1.5 | 0.7 | < 1 | 1 | < 1 | < 5 | 2.8 | 0.6 | 0.5 | 1.1 | 0.9 | < 5 | 6.6276 | < 1 | 0.5 | 0.59 | < 1 | 0.43228 | 0.8 | 0.454 | 0.4 | < 1 | 0.5 | 0.5 |  |  |  |  | 4
| 1 | 3 | < 1 | 0.52 | < 1 | < 1.5 | 0.8 | < 1 | 1 | < 1 | < 5 | 2.6 | 0.8 | 0.5 | 0.8 | 0.8 | < 5 | < 5 | < 1 | 0.5 | 2.36 | < 1 | 0.45609 | 0.9 | 0.443 | 2.3 | < 1 | 0.5 | 0.4 |  |  |  |  | 16
| 1 | 4 | < 1 | 0.46 | < 1 | < 1.5 | 0.7 | 1 | 1 | < 1 | < 5 | 1.8 | 0.6 | 0.5 | 0.4 | 1.2 | < 5 | < 5 | < 1 | 0.5 | 3.44 | < 1 | 0.54174 | 0.9 | 0.466 | 0.6 | < 1 | 0.5 | 0.5 |  |  |  |  | 0.65583
| 1 | 5 | < 1 | 0.47 | < 1 | < 1.5 | 0.7 | 2 | 1 | < 1 | < 5 | 2.8 | 0.9 | 0.5 | 1 | 1 | < 5 | 6.3998 | < 1 | 0.5 | 2.04 | < 1 | 0.44716 | 0.8 | 0.35 | 2.1 | 1 | 0.5 | 0.4 |  |  |  |  | 12
| 1 | 6 | < 1 | 0.44 | < 1 | < 1.5 | 0.7 | 6 | 1 | < 1 | < 5 | 2.7 | 1 | 0.5 | 0.6 | 0.6 | < 5 | < 5 | < 1 | 0.6 | 0.81 | < 1 | 0.39837 | 0.7 | 0.457 | 2 | < 1 | 0.5 | 0.4
| Mean |  |  | 0.47167 |  |  | 0.7 | 2.5 | 1 |  |  | 2.45 | 0.78333 | 0.51667 | 0.78333 | 0.88333 |  | 6.5137 |  | 0.51667 | 2.0533 |  | 0.45673 | 0.81667 | 0.44083 | 1.6667 | 1 | 0.5 | 0.45
| Median |  |  | 0.465 |  |  | 0.7 | 1.5 | 1 |  |  | 2.65 | 0.8 | 0.5 | 0.8 | 0.85 |  | 6.5137 |  | 0.5 | 2.2 |  | 0.45162 | 0.8 | 0.4555 | 2.05 | 1 | 0.5 | 0.45
| Std Dev. |  |  | 0.029269 |  |  | 0.063246 | 2.3805 | 0 |  |  | 0.43704 | 0.16021 | 0.040825 | 0.25626 | 0.20412 |  | 0.16108 |  | 0.040825 | 1.1626 |  | 0.047721 | 0.075277 | 0.045806 | 0.9288 |  | 0 | 0.054772
| Rel.Std.Dev. |  |  | 6.205% |  |  | 9.035% | 95.22% | 0% |  |  | 17.84% | 20.45% | 7.902% | 32.71% | 23.11% |  | 2.473% |  | 7.902% | 56.62% |  | 10.45% | 9.218% | 10.39% | 55.73% |  | 0% | 12.17%
| PDM3 |  |  | -28.08% |  |  | 6.736% | 281.2% | 52.48% |  |  | 273.6% | 19.44% | -21.22% | 19.44% | 34.69% |  | 893.2% |  | -21.22% | 213.1% |  | -30.36% | 24.52% | -32.78% | 154.1% | 52.48% | -23.76% | -31.38%
| Z-Score (Absolute) |  | 0.67 | 0.77 | 0.67 | 0.17 | 0 | 3.82 | 1.01 | 0.67 | 6.07 | 5.9 | 0.28 | 0.62 | 0.28 | 0.62 | 6.07 | 10.58 | 0.67 | 0.62 | 4.56 | 0.67 | 0.82 | 0.39 | 0.87 | 3.26 | 0.39 | 0.67 | 0.84
| Analytical results for Si in OREAS 234b (Indicative Value 0.093 wt.%) | [32 empty] | Indicative
Si | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.25 | 1.27 | [29 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | [29 empty] | 3
|  |  | 0.5g | 25g | [29 empty] | 3
| 1 | 1 | 0.11041 | 0.077 | [29 empty] | 1
| 1 | 2 | 0.13913 | 0.053 | [29 empty] | 7
| 1 | 3 | 0.11504 | 0.072 | [29 empty] | 16
| 1 | 4 | 0.10805 | 0.0625 | [29 empty] | 0.092962
| 1 | 5 | 0.12144 | 0.071 | [29 empty] | 13
| 1 | 6 | 0.13298 | 0.053
| Mean |  | 0.12117 | 0.06475
| Median |  | 0.11824 | 0.06675
| Std Dev. |  | 0.012552 | 0.010226
| Rel.Std.Dev. |  | 10.36% | 15.79%
| PDM3 |  | 30.35% | -30.35%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sm in OREAS 234b (Certified Value 1.52 ppm) | [32 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 1.572 | 1.465 | 1.3474 | 1.7 | 1.71 | 1.5 | [25 empty] | 1
| 1 | 2 | 1.585 | 1.416 | 1.3734 | 1.4 | 1.68 | 1.6 | [25 empty] | 19
| 1 | 3 | 1.572 | 1.553 | 1.3689 | 1.5 | 1.64 | 1.5 | [25 empty] | 16
| 1 | 4 | 1.54 | 1.485 | 1.3418 | 1.7 | 1.78 | 1.6 | [25 empty] | 1.5174
| 1 | 5 | 1.574 | 1.456 | 1.3375 | 1.5 | 1.68 | 1.4 | [25 empty] | 115
| 1 | 6 | 1.574 | 1.477 | 1.3434 | 1.4 | 1.66 | 1.4
| Mean |  | 1.5695 | 1.4753 | 1.3521 | 1.5333 | 1.6917 | 1.5
| Median |  | 1.573 | 1.471 | 1.3454 | 1.5 | 1.68 | 1.5
| Std Dev. |  | 0.015255 | 0.04501 | 0.015197 | 0.13663 | 0.04916 | 0.089443
| Rel.Std.Dev. |  | 0.9719% | 3.051% | 1.124% | 8.91% | 2.906% | 5.963%
| PDM3 |  | 3.435% | -2.771% | -10.89% | 1.052% | 11.49% | -1.145%
| Z-Score (Absolute) |  | 0.76 | 0.59 | 2.36 | 0.24 | 2.51 | 0.24
| Analytical results for Sn in OREAS 234b (Certified Value 0.65 ppm) | [32 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.7 | 0.65 | < 1 | 0.5083 | 0.58683 | 0.75 | 0.61 | 0.9 | 0.69 | < 10 | 0.6 | 0.6 | 0.7 | 0.6 | 0.6 | < 10 | < 5 | 0.7 | 0.6 | 44.69 | 0.5 | 0.66124 | 0.6 | 0.9 | 0.7 | 0.7 | 0.4 |  |  |  |  | 1
| 1 | 2 | 0.8 | 0.64 | < 1 | 0.56598 | 0.58338 | 0.68 | 0.58 | 0.8 | 0.67 | < 10 | 0.6 | 0.6 | 0.7 | 0.7 | 0.6 | < 10 | < 5 | 0.9 | 0.6 | 41.24 | < 0.5 | 0.70837 | 0.6 | 0.9 | 0.6 | 0.7 | 0.4 |  |  |  |  | 20
| 1 | 3 | 0.7 | 0.64 | < 1 | 0.53004 | 0.66109 | 0.85 | 0.58 | 0.9 | 0.66 | < 10 | 0.6 | 0.7 | 0.7 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.6 | 43.55 | < 0.5 | 0.69144 | 0.6 | 0.8 | 0.6 | 0.7 | 0.3 |  |  |  |  | 16
| 1 | 4 | 0.7 | 0.66 | < 1 | 0.51794 | 0.64937 | 0.71 | 0.61 | 0.9 | 0.72 | < 10 | 0.6 | 0.7 | 0.7 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.6 | 52.25 | < 0.5 | 0.70126 | 0.6 | 0.8 | 0.7 | 0.7 | 0.4 |  |  |  |  | 0.65094
| 1 | 5 | 0.7 | 0.67 | < 1 | 0.5349 | 0.59763 | 0.73 | 0.54 | 0.8 | 0.71 | < 10 | 0.6 | 0.6 | 0.7 | 0.6 | 0.6 | < 10 | < 5 | 0.8 | 0.6 | 44.08 | < 0.5 | 0.72877 | 0.6 | 0.7 | 0.7 | 0.7 | 0.4 |  |  |  |  | 116
| 1 | 6 | 0.7 | 0.65 | < 1 | 0.52686 | 0.63437 | 0.74 | 0.57 | 0.8 | 0.7 | < 10 | 0.6 | 0.6 | 0.7 | 0.6 | 0.6 | < 10 | < 5 | 0.9 | 0.6 | 43.35 | < 0.5 | 0.69254 | 0.6 | 0.7 | 0.7 | 0.7 | 0.3
| Mean |  | 0.71667 | 0.65167 |  | 0.53067 | 0.61878 | 0.74333 | 0.58167 | 0.85 | 0.69167 |  | 0.6 | 0.63333 | 0.7 | 0.65 | 0.63333 |  |  | 0.81667 | 0.6 | 44.86 | 0.5 | 0.69727 | 0.6 | 0.8 | 0.66667 | 0.7 | 0.36667
| Median |  | 0.7 | 0.65 |  | 0.52845 | 0.616 | 0.735 | 0.58 | 0.85 | 0.695 |  | 0.6 | 0.6 | 0.7 | 0.65 | 0.6 |  |  | 0.8 | 0.6 | 43.815 | 0.5 | 0.6969 | 0.6 | 0.8 | 0.7 | 0.7 | 0.4
| Std Dev. |  | 0.040825 | 0.01169 |  | 0.019709 | 0.033737 | 0.05785 | 0.026394 | 0.054772 | 0.023166 |  | 0 | 0.05164 | 1.2162e-16 | 0.054772 | 0.05164 |  |  | 0.075277 | 0 | 3.8039 |  | 0.022298 | 0 | 0.089443 | 0.05164 | 1.2162e-16 | 0.05164
| Rel.Std.Dev. |  | 5.696% | 1.794% |  | 3.714% | 5.452% | 7.783% | 4.538% | 6.444% | 3.349% |  | 0% | 8.154% | 1.737e-14% | 8.427% | 8.154% |  |  | 9.218% | 0% | 8.48% |  | 3.198% | 0% | 11.18% | 7.746% | 1.737e-14% | 14.08%
| PDM3 |  | 10.1% | 0.1123% |  | -18.48% | -4.94% | 14.19% | -10.64% | 30.58% | 6.257% |  | -7.825% | -2.704% | 7.537% | -0.1438% | -2.704% |  |  | 25.46% | -7.825% | 6792% | -23.19% | 7.118% | -7.825% | 22.9% | 2.417% | 7.537% | -43.67%
| Z-Score (Absolute) |  | 0.51 | 0.15 | 1.69 | 1.38 | 0.48 | 0.78 | 0.86 | 1.85 | 0.25 | 43.83 | 0.67 | 0.34 | 0.34 | 0.17 | 0.34 | 43.83 | 18.54 | 1.52 | 0.67 | 447 | 3.79 | 0.31 | 0.67 | 1.35 | 0 | 0.34 | 3.03
| Analytical results for Sr in OREAS 234b (Certified Value 36.9 ppm) | [32 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 22.3 | 37.71 | 143.31 | 47.9859 | 37.4518 | 55 | 47 | 48.3 | 46.5 | 39 | 37 | 39.1 | 37.5 | 32.6 | 36.8 | 35 | 48.0407 | 39.5 | 28.5 | 35.2 | 26.08 | 38.5671 | 27 | 37 | 29.3 | 33.2 | 37.6 | 27.3 |  |  |  | 1
| 1 | 2 | 24.2 | 36.72 | 164.08 | 48.4114 | 37.0073 | 52.5 | 46 | 44.6 | 45.75 | 40 | 37 | 38.2 | 38.3 | 35.9 | 35.6 | 37 | 46.2561 | 40.4 | 32.4 | 34.7 | 29.78 | 36.2328 | 26 | 37.8 | 31 | 33.9 | 37.7 | 27.1 |  |  |  | 21
| 1 | 3 | 23.8 | 36.78 | 154.47 | 46.9462 | 37.5539 | 54 | 45 | 46.7 | 47.6 | 40 | 37 | 39.4 | 37.8 | 36.3 | 36.9 | 36 | 46.57 | 39.7 | 33 | 36.6 | 29.9 | 34.4286 | 27 | 36.8 | 31.2 | 34 | 37.9 | 26.9 |  |  |  | 16
| 1 | 4 | 25.3 | 37.22 | 154.02 | 45.5454 | 37.3899 | 49.8 | 45 | 47.3 | 45.76 | 40 | 37 | 38 | 40.6 | 35.6 | 36.7 | 33 | 44.3173 | 43 | 30.1 | 36 | 29.41 | 36.0933 | 27 | 36.9 | 31.1 | 32.4 | 36.9 | 27.7 |  |  |  | 36.9431
| 1 | 5 | 26.7 | 36.79 | 129.82 | 47.4743 | 38.6171 | 55.2 | 39 | 48.5 | 48.88 | 42 | 37 | 38.4 | 38 | 33.9 | 36.8 | 34 | 47.3294 | 44.2 | 29.5 | 36.4 | 28.57 | 37.8975 | 26 | 36.8 | 29.7 | 33.7 | 38.3 | 27.8 |  |  |  | 117
| 1 | 6 | 23.6 | 37.34 | 154.58 | 48.2782 | 38.3994 | 53.1 | 44 | 48.7 | 48.1 | 41 | 37 | 35.1 | 36.5 | 34.6 | 36.2 | 37 | 48.6318 | 45.7 | 33.2 | 36 | 27.5 | 38.6694 | 25 | 37 | 29.6 | 33.8 | 37.3 | 26.7
| Mean |  | 24.3167 | 37.0933 | 150.0467 | 47.4402 | 37.7366 | 53.2667 | 44.3333 | 47.35 | 47.0983 | 40.3333 | 37 | 38.0333 | 38.1167 | 34.8167 | 36.5 | 35.3333 | 46.8575 | 42.0833 | 31.1167 | 35.8167 | 28.54 | 36.9815 | 26.3333 | 37.05 | 30.3167 | 33.5 | 37.6167 | 27.25
| Median |  | 24 | 37.005 | 154.245 | 47.7301 | 37.5029 | 53.55 | 45 | 47.8 | 47.05 | 40 | 37 | 38.3 | 37.9 | 35.1 | 36.75 | 35.5 | 46.9497 | 41.7 | 31.25 | 36 | 28.99 | 37.0652 | 26.5 | 36.95 | 30.35 | 33.75 | 37.65 | 27.2
| Std Dev. |  | 1.5171 | 0.39667 | 11.894 | 1.0751 | 0.62946 | 1.9957 | 2.8048 | 1.5515 | 1.2959 | 1.0328 | 0 | 1.5345 | 1.3644 | 1.3992 | 0.50596 | 1.633 | 1.5276 | 2.5918 | 2.0014 | 0.72778 | 1.5019 | 1.6774 | 0.8165 | 0.37815 | 0.87044 | 0.60663 | 0.48339 | 0.43704
| Rel.Std.Dev. |  | 6.239% | 1.069% | 7.927% | 2.266% | 1.668% | 3.747% | 6.327% | 3.277% | 2.751% | 2.561% | 0% | 4.035% | 3.58% | 4.019% | 1.386% | 4.622% | 3.26% | 6.159% | 6.432% | 2.032% | 5.262% | 4.536% | 3.101% | 1.021% | 2.871% | 1.811% | 1.285% | 1.604%
| PDM3 |  | -34.18% | 0.4066% | 306.2% | 28.41% | 2.148% | 44.19% | 20% | 28.17% | 27.49% | 9.177% | 0.154% | 2.951% | 3.177% | -5.756% | -1.199% | -4.357% | 26.84% | 13.91% | -15.77% | -3.049% | -22.75% | 0.1038% | -28.72% | 0.2893% | -17.94% | -9.32% | 1.823% | -26.24%
| Z-Score (Absolute) |  | 2 | 0 | 17.75 | 1.63 | 0.1 | 2.54 | 1.14 | 1.61 | 1.58 | 0.51 | 0.01 | 0.15 | 0.16 | 0.35 | 0.09 | 0.27 | 1.54 | 0.79 | 0.94 | 0.2 | 1.34 | 0.01 | 1.69 | 0 | 1.06 | 0.56 | 0.09 | 1.54
| Analytical results for Ta in OREAS 234b (Certified Value < 0.01 ppm) | [32 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.51 | 0.14 | < 0.01 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.17 | 0.1 | < 0.01 | < 0.01 | [11 empty] | 22
| 1 | 3 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.74 | 0.07 | < 0.01 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.21 | 0.09 | 0.03 | < 0.01 | [11 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.01 | 0.01478 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.21 | 0.11 | < 0.01 | < 0.01 | [11 empty] | 118
| 1 | 6 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.22 | 0.06 | < 0.01 | < 0.01
| Mean |  |  |  | 0.01478 |  | 0.033333 |  |  | 5 |  | 0.01 | 0.01 |  |  |  |  | 0.01 | 0.34333 | 0.095 | 0.03
| Median |  |  |  | 0.01478 |  | 0.03 |  |  | 5 |  | 0.01 | 0.01 |  |  |  |  | 0.01 | 0.215 | 0.095 | 0.03
| Std Dev. | [5 empty] | 0.005164 |  |  | 0 |  | 0 | 0 |  |  |  |  | 0 | 0.23062 | 0.02881
| Rel.Std.Dev. | [5 empty] | 15.49% |  |  | 0% |  | 0% | 0% |  |  |  |  | 0% | 67.17% | 30.33%
PDM3
| Z-Score (Absolute) |  | 2.02 | 0.67 | 0.45 | 2.02 | 3.15 | 2.02 | 0.67 | NA | 5.39 | 0.45 | 0 | 0.67 | 0.67 | 2.02 | 32.37 | 0.45 | 44.95 | 11.46 | 0.11 | 0.67
| Analytical results for Tb in OREAS 234b (Certified Value 0.3 ppm) | [32 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 3
| 1 | 1 | 0.29 | 0.303 | 0.297 | 0.31316 | 0.4 | 0.31 | 0.3 | 0.31 | [23 empty] | 1
| 1 | 2 | 0.29 | 0.309 | 0.291 | 0.30585 | 0.3 | 0.3 | 0.25 | 0.32 | [23 empty] | 23
| 1 | 3 | 0.29 | 0.306 | 0.329 | 0.30627 | 0.4 | 0.3 | 0.25 | 0.32 | [23 empty] | 16
| 1 | 4 | 0.31 | 0.306 | 0.304 | 0.29927 | 0.3 | 0.31 | 0.25 | 0.31 | [23 empty] | 0.30411
| 1 | 5 | 0.3 | 0.308 | 0.303 | 0.29798 | 0.3 | 0.28 | 0.25 | 0.31 | [23 empty] | 119
| 1 | 6 | 0.3 | 0.3 | 0.314 | 0.29555 | 0.4 | 0.29 | 0.25 | 0.32
| Mean |  | 0.29667 | 0.30533 | 0.30633 | 0.30302 | 0.35 | 0.29833 | 0.25833 | 0.315
| Median |  | 0.295 | 0.306 | 0.3035 | 0.30256 | 0.35 | 0.3 | 0.25 | 0.315
| Std Dev. |  | 0.008165 | 0.0033267 | 0.013501 | 0.0065827 | 0.054772 | 0.01169 | 0.020412 | 0.0054772
| Rel.Std.Dev. |  | 2.752% | 1.09% | 4.407% | 2.172% | 15.65% | 3.919% | 7.902% | 1.739%
| PDM3 |  | -2.449% | 0.401% | 0.7299% | -0.3611% | 15.09% | -1.901% | -15.05% | 3.58%
| Z-Score (Absolute) |  | 0.91 | 0.33 | 0.48 | 0 | NA | 0.67 | 6.43 | 1.73
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 234b (Certified Value 0.1 ppm) | [32 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 0.09 | 0.1 | 0.1097 | 0.25 | 0.1 | 0.05 | < 0.1 | < 5 | < 0.1 | 0.09 | 0.1 | 0.07 | 0.08 | < 5 | 0.11 | 0.08 | < 0.1 | 0.1467 | 0.1 | 0.103 | 0.11 | 0.09 | 0.09 | 0.1 | [7 empty] | 1
| 1 | 2 | 0.16 | 0.1 | 0.11787 | 0.13 | 0.09 | 0.05 | < 0.1 | < 5 | < 0.1 | 0.08 | 0.11 | 0.11 | 0.08 | < 5 | 0.12 | 0.08 | < 0.1 | 0.10388 | 0.1 | 0.077 | < 0.01 | 0.1 | 0.09 | 0.12 | [7 empty] | 24
| 1 | 3 | 0.1 | 0.09 | 0.13441 | 0.11 | 0.1 | 0.06 | < 0.1 | < 5 | < 0.1 | 0.09 | 0.1 | 0.09 | 0.1 | < 5 | 0.1 | 0.1 | < 0.1 | 0.11676 | 0.11 | 0.309 | 0.02 | 0.11 | 0.09 | 0.11 | [7 empty] | 16
| 1 | 4 | 0.12 | 0.1 | 0.13181 | 0.08 | 0.11 | 0.06 | < 0.1 | < 5 | < 0.1 | 0.07 | 0.1 | 0.09 | 0.1 | < 5 | 0.09 | 0.09 | < 0.1 | 0.1019 | 0.1 | < 0.03 | 0.03 | 0.09 | 0.1 | 0.08 | [7 empty] | 0.10124
| 1 | 5 | 0.11 | 0.1 | 0.11651 | 0.08 | 0.11 | < 0.05 | 0.1 | < 5 | 0.1 | 0.09 | 0.1 | 0.11 | 0.09 | < 5 | 0.12 | 0.11 | < 0.1 | 0.15248 | 0.11 | 0.257 | < 0.01 | 0.1 | 0.08 | 0.12 | [7 empty] | 120
| 1 | 6 | 0.08 | 0.1 | 0.13116 | 0.1 | 0.1 | 0.07 | < 0.1 | < 5 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | < 5 | 0.1 | 0.08 | < 0.1 | 0.09615 | 0.1 | 0.077 | < 0.01 | 0.11 | 0.09 | 0.15
| Mean |  | 0.11 | 0.098333 | 0.12358 | 0.125 | 0.10167 | 0.058 | 0.1 |  | 0.1 | 0.086667 | 0.10167 | 0.093333 | 0.091667 |  | 0.10667 | 0.09 |  | 0.11965 | 0.10333 | 0.1646 | 0.053333 | 0.1 | 0.09 | 0.11333
| Median |  | 0.105 | 0.1 | 0.12452 | 0.105 | 0.1 | 0.06 | 0.1 |  | 0.1 | 0.09 | 0.1 | 0.09 | 0.095 |  | 0.105 | 0.085 |  | 0.11032 | 0.1 | 0.103 | 0.03 | 0.1 | 0.09 | 0.115
| Std Dev. |  | 0.028284 | 0.0040825 | 0.010176 | 0.064109 | 0.0075277 | 0.0083666 |  |  | 0 | 0.010328 | 0.0040825 | 0.015055 | 0.0098319 |  | 0.012111 | 0.012649 |  | 0.024223 | 0.005164 | 0.11015 | 0.049329 | 0.0089443 | 0.0063246 | 0.023381
| Rel.Std.Dev. |  | 25.71% | 4.152% | 8.234% | 51.29% | 7.404% | 14.43% |  |  | 0% | 11.92% | 4.016% | 16.13% | 10.73% |  | 11.35% | 14.05% |  | 20.25% | 4.997% | 66.92% | 92.49% | 8.944% | 7.027% | 20.63%
| PDM3 |  | 8.65% | -2.874% | 22.06% | 23.47% | 0.4185% | -42.71% | -1.228% |  | -1.228% | -14.4% | 0.4185% | -7.813% | -9.459% |  | 5.357% | -11.1% |  | 18.18% | 2.065% | 62.58% | -47.32% | -1.228% | -11.1% | 11.94%
| Z-Score (Absolute) |  | 0.48 | 0.19 | 1.27 | 1.35 | 0 | 2.84 | NA | 138.62 | NA | 0.87 | 0 | 0.48 | 0.58 | 138.62 | 0.29 | 0.67 | 2.99 | 1.04 | 0.1 | 2.2 | 4.19 | 0.1 | 0.67 | 0.67
Results from laboratories 1.11 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 234b (Certified Value 0.8 ppm) | [32 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.5 | 0.824 | 1.252 | 0.6031 | 0.81113 | 0.9 | 0.84 | 0.9 | 0.79 | 0.8 | 0.8 | 0.8 | 0.7 | < 100 | 0.8 | 0.76 | 0.6 | 0.8 | 0.8 | 0.8 | 1.4 | 0.8 | 0.8 | 0.8 | [7 empty] | 1
| 1 | 2 | 1 | 0.822 | 0.665 | 0.5951 | 0.81914 | 0.8 | 0.83 | 0.96 | 0.79 | 0.9 | 0.8 | 0.8 | 0.8 | < 100 | 0.8 | 0.75 | 0.7 | 0.8 | 0.8 | 0.8 | 1.3 | 0.8 | 0.8 | 0.8 | [7 empty] | 25
| 1 | 3 | 0.7 | 0.82 | 0.739 | 0.58837 | 0.82672 | 0.8 | 0.88 | 0.96 | 0.78 | 0.9 | 0.8 | 0.8 | 0.8 | < 100 | 0.8 | 0.85 | 0.7 | 0.8 | 0.8 | 0.8 | 1.3 | 0.7 | 0.8 | 0.8 | [7 empty] | 16
| 1 | 4 | 0.6 | 0.822 | 0.684 | 0.55833 | 0.80243 | 0.8 | 0.85 | 0.93 | 0.81 | 0.9 | 0.8 | 0.8 | 0.8 | < 100 | 0.7 | 0.79 | 0.6 | 0.8 | 0.8 | 0.8 | 1.2 | 0.8 | 0.8 | 0.8 | [7 empty] | 0.80163
| 1 | 5 | 0.5 | 0.824 | 0.68 | 0.58914 | 0.84659 | 0.8 | 0.86 | 0.96 | 0.79 | 0.8 | 0.8 | 0.8 | 0.8 | < 100 | 0.8 | 0.84 | 0.6 | 0.8 | 0.8 | 0.8 | 1.2 | 0.7 | 0.8 | 0.8 | [7 empty] | 121
| 1 | 6 | 0.5 | 0.824 | 0.718 | 0.58168 | 0.81387 | 0.8 | 0.84 | 0.94 | 0.77 | 0.8 | 0.7 | 0.7 | 0.8 | < 100 | 0.8 | 0.79 | 0.7 | 0.8 | 0.7 | 0.8 | 1.2 | 0.8 | 0.8 | 0.8
| Mean |  | 0.63333 | 0.82267 | 0.78967 | 0.58595 | 0.81998 | 0.81667 | 0.85 | 0.94167 | 0.78833 | 0.85 | 0.78333 | 0.78333 | 0.78333 |  | 0.78333 | 0.79667 | 0.65 | 0.8 | 0.78333 | 0.8 | 1.2667 | 0.76667 | 0.8 | 0.8
| Median |  | 0.55 | 0.823 | 0.701 | 0.58876 | 0.81651 | 0.8 | 0.845 | 0.95 | 0.79 | 0.85 | 0.8 | 0.8 | 0.8 |  | 0.8 | 0.79 | 0.65 | 0.8 | 0.8 | 0.8 | 1.25 | 0.8 | 0.8 | 0.8
| Std Dev. |  | 0.19664 | 0.001633 | 0.22812 | 0.01533 | 0.015351 | 0.040825 | 0.017889 | 0.024014 | 0.013292 | 0.054772 | 0.040825 | 0.040825 | 0.040825 |  | 0.040825 | 0.040825 | 0.054772 | 1.2162e-16 | 0.040825 | 1.2162e-16 | 0.08165 | 0.05164 | 1.2162e-16 | 1.2162e-16
| Rel.Std.Dev. |  | 31.05% | 0.1985% | 28.89% | 2.616% | 1.872% | 4.999% | 2.105% | 2.55% | 1.686% | 6.444% | 5.212% | 5.212% | 5.212% |  | 5.212% | 5.124% | 8.427% | 1.52e-14% | 5.212% | 1.52e-14% | 6.446% | 6.736% | 1.52e-14% | 1.52e-14%
| PDM3 |  | -20.99% | 2.625% | -1.492% | -26.9% | 2.29% | 1.876% | 6.034% | 17.47% | -1.658% | 6.034% | -2.282% | -2.282% | -2.282% |  | -2.282% | -0.6187% | -18.91% | -0.2029% | -2.282% | -0.2029% | 58.01% | -4.361% | -0.2029% | -0.2029%
| Z-Score (Absolute) |  | 6.68 | 0.98 | 0.35 | 8.59 | 0.88 | 0.74 | 2.09 | 5.8 | 0.4 | 2.09 | 0.61 | 0.61 | 0.61 | 1990.63 | 0.61 | 0.07 | 6 | 0.07 | 0.61 | 0.07 | 18.95 | 1.28 | 0.07 | 0.07
| Analytical results for Ti in OREAS 234b (Certified Value 0.315 wt.%) | [32 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.31 | 0.272 | 0.34501 | 0.4262 | 0.32694 | 0.38 | 0.4232 | 0.434 | 0.3768 | 0.32 | 0.249 | 0.31 | 0.332 | 0.291 | < 0.01 | 0.286 | 0.25 | 0.45126 | 0.3643 | 0.212 | 0.26 | 0.18213 | 0.32222 | 0.1135 | 0.292 | 0.223 | 0.3 | 0.276 | 0.209 |  |  | 1
| 1 | 2 | 0.3 | 0.2789 | 0.34077 | 0.42891 | 0.32194 | 0.37 | 0.413 | 0.425 | 0.369 | 0.32 | 0.255 | 0.313 | 0.297 | 0.305 | < 0.01 | 0.287 | 0.27 | 0.43599 | 0.3679 | 0.263 | 0.26 | 0.24599 | 0.30537 | 0.1105 | 0.291 | 0.25 | 0.29 | 0.29 | 0.191 |  |  | 26
| 1 | 3 | 0.31 | 0.2732 | 0.35512 | 0.43143 | 0.32907 | 0.37 | 0.4005 | 0.43 | 0.3746 | 0.31 | 0.256 | 0.316 | 0.31 | 0.301 | < 0.01 | 0.294 | 0.26 | 0.42615 | 0.3569 | 0.254 | 0.29 | 0.25389 | 0.29088 | 0.1125 | 0.29 | 0.261 | 0.29 | 0.287 | 0.2 |  |  | 16
| 1 | 4 | 0.32 | 0.2703 | 0.35897 | 0.42877 | 0.32451 | 0.37 | 0.4094 | 0.435 | 0.3804 | 0.31 | 0.262 | 0.307 | 0.304 | 0.298 | < 0.01 | 0.293 | 0.23 | 0.42944 | 0.3516 | 0.201 | 0.28 | 0.2426 | 0.30358 | 0.1115 | 0.291 | 0.252 | 0.3 | 0.285 | 0.2 |  |  | 0.31453
| 1 | 5 | 0.33 | 0.2795 | 0.36142 | 0.42918 | 0.33019 | 0.38 | 0.3501 | 0.43 | 0.3823 | 0.3 | 0.262 | 0.312 | 0.31 | 0.286 | < 0.01 | 0.304 | 0.25 | 0.45758 | 0.3525 | 0.208 | 0.29 | 0.23426 | 0.32613 | 0.103 | 0.279 | 0.239 | 0.31 | 0.3 | 0.209 |  |  | 122
| 1 | 6 | 0.31 | 0.2813 | 0.33186 | 0.4283 | 0.32915 | 0.39 | 0.3933 | 0.433 | 0.3691 | 0.31 | 0.261 | 0.289 | 0.299 | 0.298 | < 0.01 | 0.289 | 0.27 | 0.44908 | 0.3726 | 0.26 | 0.28 | 0.21009 | 0.33001 | 0.1025 | 0.285 | 0.225 | 0.3 | 0.294 | 0.206
| Mean |  | 0.31333 | 0.27587 | 0.34886 | 0.4288 | 0.32697 | 0.37667 | 0.39825 | 0.43117 | 0.37537 | 0.31167 | 0.2575 | 0.30783 | 0.30867 | 0.2965 |  | 0.29217 | 0.255 | 0.44158 | 0.36097 | 0.233 | 0.27667 | 0.22816 | 0.31303 | 0.10892 | 0.288 | 0.24167 | 0.29833 | 0.28867 | 0.2025
| Median |  | 0.31 | 0.27605 | 0.35007 | 0.42884 | 0.32801 | 0.375 | 0.40495 | 0.4315 | 0.3757 | 0.31 | 0.2585 | 0.311 | 0.307 | 0.298 |  | 0.291 | 0.255 | 0.44254 | 0.3606 | 0.233 | 0.28 | 0.23843 | 0.31379 | 0.111 | 0.2905 | 0.2445 | 0.3 | 0.2885 | 0.203
| Std Dev. |  | 0.010328 | 0.004582 | 0.011564 | 0.0016783 | 0.0031887 | 0.008165 | 0.025739 | 0.003656 | 0.0055831 | 0.0075277 | 0.0051672 | 0.009704 | 0.012644 | 0.006892 |  | 0.0066156 | 0.015166 | 0.012826 | 0.0086078 | 0.028844 | 0.013663 | 0.027095 | 0.015383 | 0.0048828 | 0.0050596 | 0.015384 | 0.0075277 | 0.0081894 | 0.0069498
| Rel.Std.Dev. |  | 3.296% | 1.661% | 3.315% | 0.3914% | 0.9752% | 2.168% | 6.463% | 0.8479% | 1.487% | 2.415% | 2.007% | 3.152% | 4.096% | 2.324% |  | 2.264% | 5.947% | 2.904% | 2.385% | 12.38% | 4.938% | 11.88% | 4.914% | 4.483% | 1.757% | 6.366% | 2.523% | 2.837% | 3.432%
| PDM3 |  | -0.3814% | -12.29% | 10.91% | 36.33% | 3.954% | 19.75% | 26.62% | 37.08% | 19.34% | -0.9113% | -18.13% | -2.13% | -1.865% | -5.733% |  | -7.111% | -18.93% | 40.39% | 14.76% | -25.92% | -12.04% | -27.46% | -0.4772% | -65.37% | -8.436% | -23.17% | -5.15% | -8.224% | -35.62%
| Z-Score (Absolute) |  | 0.23 | 0.35 | 0.79 | 2.03 | 0.45 | 1.22 | 1.55 | 2.07 | 1.2 | 0.21 | 0.64 | 0.15 | 0.16 | 0.03 | 4.56 | 0.1 | 0.67 | 2.23 | 0.97 | 1.02 | 0.34 | 1.09 | 0.23 | 2.95 | 0.16 | 0.88 | 0 | 0.15 | 1.49
| Analytical results for Tl in OREAS 234b (Certified Value 0.12 ppm) | [32 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.12 | 0.12 | 0.12037 | 0.12 | 0.11 | 0.12 | < 10 | 0.13 | 0.1 | 0.12 | 0.12 | 0.12 | < 5 | 0.11 | 0.11 | 0.12 | 6.4 | 0.1 | 0.12413 | 0.11 | 0.123 | 0.04 | 0.11 | 0.12 | 0.1 | [6 empty] | 1
| 1 | 2 | 0.13 | 0.13 | 0.11966 | 0.12 | 0.13 | 0.11 | < 10 | 0.13 | 0.11 | 0.12 | 0.12 | 0.11 | < 5 | 0.1 | 0.12 | 0.12 | 3.62 | 0.1 | 0.1236 | 0.11 | 0.128 | 0.03 | 0.12 | 0.12 | 0.11 | [6 empty] | 27
| 1 | 3 | 0.12 | 0.12 | 0.12144 | 0.12 | 0.13 | 0.12 | < 10 | 0.13 | 0.11 | 0.12 | 0.13 | 0.11 | < 5 | 0.11 | 0.12 | 0.12 | 4.57 | 0.1 | 0.12435 | 0.12 | 0.12 | 0.04 | 0.11 | 0.12 | 0.11 | [6 empty] | 16
| 1 | 4 | 0.12 | 0.12 | 0.11884 | 0.11 | 0.12 | 0.12 | < 10 | 0.13 | 0.11 | 0.12 | 0.12 | 0.11 | < 5 | 0.11 | 0.11 | 0.12 | 6.33 | 0.1 | 0.12888 | 0.12 | 0.127 | 0.03 | 0.11 | 0.12 | 0.11 | [6 empty] | 0.11831
| 1 | 5 | 0.12 | 0.13 | 0.12988 | 0.11 | 0.12 | 0.12 | < 10 | 0.13 | 0.11 | 0.11 | 0.12 | 0.12 | < 5 | 0.12 | 0.11 | 0.12 | 3.4 | 0.1 | 0.12482 | 0.12 | 0.122 | 0.03 | 0.12 | 0.12 | 0.11 | [6 empty] | 123
| 1 | 6 | 0.12 | 0.12 | 0.12923 | 0.12 | 0.13 | 0.11 | < 10 | 0.12 | 0.14 | 0.11 | 0.11 | 0.12 | < 5 | 0.12 | 0.12 | 0.12 | 4.83 | 0.1 | 0.12587 | 0.11 | 0.132 | 0.03 | 0.11 | 0.12 | 0.11
| Mean |  | 0.12167 | 0.12333 | 0.12324 | 0.11667 | 0.12333 | 0.11667 |  | 0.12833 | 0.11333 | 0.11667 | 0.12 | 0.115 |  | 0.11167 | 0.115 | 0.12 | 4.8583 | 0.1 | 0.12528 | 0.115 | 0.12533 | 0.033333 | 0.11333 | 0.12 | 0.10833
| Median |  | 0.12 | 0.12 | 0.1209 | 0.12 | 0.125 | 0.12 |  | 0.13 | 0.11 | 0.12 | 0.12 | 0.115 |  | 0.11 | 0.115 | 0.12 | 4.7 | 0.1 | 0.12459 | 0.115 | 0.125 | 0.03 | 0.11 | 0.12 | 0.11
| Std Dev. |  | 0.0040825 | 0.005164 | 0.0049709 | 0.005164 | 0.008165 | 0.005164 |  | 0.0040825 | 0.013663 | 0.005164 | 0.0063246 | 0.0054772 |  | 0.0075277 | 0.0054772 | 0 | 1.2874 | 1.5202e-17 | 0.0019251 | 0.0054772 | 0.0044572 | 0.005164 | 0.005164 | 0 | 0.0040825
| Rel.Std.Dev. |  | 3.355% | 4.187% | 4.034% | 4.426% | 6.62% | 4.426% |  | 3.181% | 12.06% | 4.426% | 5.27% | 4.763% |  | 6.741% | 4.763% | 0% | 26.5% | 1.52e-14% | 1.537% | 4.763% | 3.556% | 15.49% | 4.556% | 0% | 3.768%
| PDM3 |  | 2.841% | 4.249% | 4.167% | -1.386% | 4.249% | -1.386% |  | 8.476% | -4.203% | -1.386% | 1.432% | -2.795% |  | -5.612% | -2.795% | 1.432% | 4007% | -15.47% | 5.891% | -2.795% | 5.94% | -71.82% | -4.203% | 1.432% | -8.43%
| Z-Score (Absolute) |  | 0.22 | 0.45 | 0.44 | 0.45 | 0.45 | 0.45 | 658.13 | 1.12 | 0.9 | 0.45 | 0 | 0.67 | 320.97 | 1.12 | 0.67 | 0 | 639.02 | NA | 0.71 | 0.67 | 0.72 | 11.69 | 0.9 | 0 | 1.57
Results from laboratory 1.27 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Tm in OREAS 234b (Certified Value 0.16 ppm) | [32 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | [26 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [26 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [26 empty] | 2
| 1 | 1 | 0.17 | 0.15 | 0.14718 | 0.2 | 0.15 | [26 empty] | 1
| 1 | 2 | 0.167 | 0.147 | 0.15187 | 0.2 | 0.15 | [26 empty] | 28
| 1 | 3 | 0.167 | 0.168 | 0.14285 | 0.1 | 0.1 | [26 empty] | 16
| 1 | 4 | 0.165 | 0.157 | 0.14483 | 0.2 | 0.15 | [26 empty] | 0.15719
| 1 | 5 | 0.17 | 0.158 | 0.14245 | 0.2 | 0.15 | [26 empty] | 124
| 1 | 6 | 0.164 | 0.163 | 0.14047 | 0.2 | 0.1
| Mean |  | 0.16717 | 0.15717 | 0.14494 | 0.18333 | 0.13333
| Median |  | 0.167 | 0.1575 | 0.14384 | 0.2 | 0.15
| Std Dev. |  | 0.0024833 | 0.0078337 | 0.0040864 | 0.040825 | 0.02582
| Rel.Std.Dev. |  | 1.486% | 4.984% | 2.819% | 22.27% | 19.36%
| PDM3 |  | 6.348% | -0.01382% | -7.791% | 16.63% | -15.18%
| Z-Score (Absolute) |  | 0.55 | 0 | 0.67 | 1.44 | 1.31
| Analytical results for U in OREAS 234b (Certified Value 0.2 ppm) | [32 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.21 | 0.217 | 0.303 | 0.17173 | 0.22351 | 0.2 | 0.28 | 0.24 | 0.22 | 0.2 | 0.2 | 0.21 | 0.18 | 1465.019 | 0.21 | < 10 | 0.19 | 0.2 | 0.2 | 0.2 | 0.2 | 0.195 | 0.17 | 0.2 | 0.21 | 0.18 | [5 empty] | 1
| 1 | 2 | 0.21 | 0.212 | 0.196 | 0.1617 | 0.22547 | 0.2 | 0.26 | 0.27 | 0.22 | 0.2 | 0.2 | 0.2 | 0.2 | 1443.755 | 0.2 | < 10 | 0.2 | 0.2 | 0.21 | 0.2 | 0.19 | 0.198 | 0.19 | 0.21 | 0.21 | 0.18 | [5 empty] | 29
| 1 | 3 | 0.21 | 0.211 | 0.23 | 0.15559 | 0.2419 | 0.2 | 0.32 | 0.27 | 0.22 | 0.2 | 0.19 | 0.2 | 0.2 | 1443.596 | 0.2 | < 10 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.192 | 0.18 | 0.2 | 0.21 | 0.18 | [5 empty] | 16
| 1 | 4 | 0.21 | 0.213 | 0.214 | 0.15693 | 0.22437 | 0.2 | 0.42 | 0.26 | 0.22 | 0.2 | 0.19 | 0.19 | 0.19 | 1490.74 | 0.2 | < 10 | 0.24 | 0.2 | 0.2 | 0.2 | 0.19 | 0.199 | 0.18 | 0.2 | 0.21 | 0.19 | [5 empty] | 0.20321
| 1 | 5 | 0.21 | 0.216 | 0.214 | 0.15727 | 0.23819 | 0.2 | 0.23 | 0.27 | 0.23 | 0.2 | 0.2 | 0.2 | 0.2 | 1498.519 | 0.2 | < 10 | 0.23 | 0.2 | 0.2 | 0.2 | 0.19 | 0.19 | 0.17 | 0.19 | 0.21 | 0.18 | [5 empty] | 125
| 1 | 6 | 0.21 | 0.215 | 0.212 | 0.16312 | 0.23341 | 0.2 | 0.24 | 0.27 | 0.23 | 0.2 | 0.18 | 0.19 | 0.19 | 1547.139 | 0.2 | < 10 | 0.23 | 0.2 | 0.21 | 0.2 | 0.19 | 0.193 | 0.17 | 0.2 | 0.21 | 0.18
| Mean |  | 0.21 | 0.214 | 0.22817 | 0.16106 | 0.23114 | 0.2 | 0.29167 | 0.26333 | 0.22333 | 0.2 | 0.19333 | 0.19833 | 0.19333 | 1481.4613 | 0.20167 |  | 0.21667 | 0.2 | 0.20333 | 0.2 | 0.19333 | 0.1945 | 0.17667 | 0.2 | 0.21 | 0.18167
| Median |  | 0.21 | 0.214 | 0.214 | 0.15948 | 0.22944 | 0.2 | 0.27 | 0.27 | 0.22 | 0.2 | 0.195 | 0.2 | 0.195 | 1477.8795 | 0.2 |  | 0.22 | 0.2 | 0.2 | 0.2 | 0.19 | 0.194 | 0.175 | 0.2 | 0.21 | 0.18
| Std Dev. |  | 0 | 0.0023664 | 0.038212 | 0.005997 | 0.0078333 | 3.0405e-17 | 0.070545 | 0.012111 | 0.005164 | 3.0405e-17 | 0.008165 | 0.0075277 | 0.008165 | 39.5386 | 0.0040825 |  | 0.019664 | 3.0405e-17 | 0.005164 | 3.0405e-17 | 0.005164 | 0.0035071 | 0.008165 | 0.0063246 | 0 | 0.0040825
| Rel.Std.Dev. |  | 0% | 1.106% | 16.75% | 3.724% | 3.389% | 1.52e-14% | 24.19% | 4.599% | 2.312% | 1.52e-14% | 4.223% | 3.795% | 4.223% | 2.669% | 2.024% |  | 9.076% | 1.52e-14% | 2.54% | 1.52e-14% | 2.671% | 1.803% | 4.622% | 3.162% | 0% | 2.247%
| PDM3 |  | 3.343% | 5.312% | 12.28% | -20.74% | 13.75% | -1.578% | 43.53% | 29.59% | 9.905% | -1.578% | -4.859% | -2.398% | -4.859% | 728943% | -0.7577% |  | 6.624% | -1.578% | 0.06247% | -1.578% | -4.859% | -4.285% | -13.06% | -1.578% | 3.343% | -10.6%
| Z-Score (Absolute) |  | 0.17 | 0.37 | 1.09 | 2.31 | 1.24 | NA | 4.3 | 2.87 | 0.84 | NA | 0.67 | 0.42 | 0.67 | 74911.73 | 0.25 | 242.41 | 0.51 | NA | 0.17 | NA | 0.67 | 0.62 | 1.52 | 0.34 | 0.17 | 1.26
Results from laboratories 1.07, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 234b (Certified Value 127 ppm) | [32 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 131 | 131.7 | 119.52 | 141.6161 | 140.9898 | 149 | 174 | 177 | 161 | 135 | 126 | 130 | 131 | 122 | < 10 | 128 | 116 | 173.0402 | 152 | 100 | 121 | 93.72 | 141.7812 | 90 | 126 | 104 | 123 | 129 | 95 |  |  | 1
| 1 | 2 | 130 | 132.4 | 118.9 | 142.7077 | 139.4107 | 140 | 163 | 172 | 159 | 140 | 130 | 130 | 123 | 125 | < 10 | 125 | 119 | 173.0366 | 152 | 120 | 119 | 111.64 | 134.6236 | 85 | 125 | 111 | 121 | 132 | 91.7 |  |  | 30
| 1 | 3 | 132 | 130.8 | 123.37 | 141.2496 | 142.1409 | 143 | 161 | 173 | 163 | 134 | 128 | 130 | 126 | 125 | < 10 | 127 | 117 | 173.616 | 143 | 118 | 126 | 113.34 | 137.3924 | 90 | 125 | 115 | 123 | 132 | 93 |  |  | 16
| 1 | 4 | 137 | 133.9 | 123.04 | 139.382 | 139.6648 | 140 | 170 | 172 | 164 | 137 | 129 | 127 | 125 | 124 | < 10 | 127 | 107 | 170.0426 | 153 | 105 | 132 | 110.3 | 133.1332 | 90 | 126 | 111 | 121 | 132 | 95.2 |  |  | 127.4601
| 1 | 5 | 139 | 132.2 | 124.56 | 141.2543 | 142.8101 | 141 | 144 | 174 | 167 | 138 | 127 | 129 | 125 | 120 | < 10 | 129 | 112 | 181.095 | 141 | 105 | 124 | 107.63 | 140.5484 | 85 | 123 | 108 | 122 | 137 | 96.2 |  |  | 126
| 1 | 6 | 134 | 131.7 | 116.28 | 138.5886 | 143.1726 | 140 | 155 | 173 | 159 | 136 | 125 | 121 | 124 | 123 | < 10 | 125 | 121 | 176.0002 | 147 | 114 | 128 | 100.87 | 142.9404 | 85 | 123 | 103 | 121 | 135 | 92.8
| Mean |  | 133.8333 | 132.1167 | 120.945 | 140.7997 | 141.3648 | 142.1667 | 161.1667 | 173.5 | 162.1667 | 136.6667 | 127.5 | 127.8333 | 125.6667 | 123.1667 |  | 126.8333 | 115.3333 | 174.4718 | 148 | 110.3333 | 125 | 106.25 | 138.4032 | 87.5 | 124.6667 | 108.6667 | 121.8333 | 132.8333 | 93.9833
| Median |  | 133 | 131.95 | 121.28 | 141.252 | 141.5653 | 140.5 | 162 | 173 | 162 | 136.5 | 127.5 | 129.5 | 125 | 123.5 |  | 127 | 116.5 | 173.3281 | 149.5 | 109.5 | 125 | 108.965 | 138.9704 | 87.5 | 125 | 109.5 | 121.5 | 132 | 94
| Std Dev. |  | 3.5449 | 1.0342 | 3.2035 | 1.5242 | 1.6005 | 3.5449 | 10.7595 | 1.8708 | 3.1252 | 2.1602 | 1.8708 | 3.5449 | 2.8048 | 1.9408 |  | 1.6021 | 5.0859 | 3.7593 | 5.1381 | 8.1158 | 4.7329 | 7.5309 | 3.992 | 2.7386 | 1.3663 | 4.5898 | 0.98319 | 2.7869 | 1.7325
| Rel.Std.Dev. |  | 2.649% | 0.7828% | 2.649% | 1.083% | 1.132% | 2.494% | 6.676% | 1.078% | 1.927% | 1.581% | 1.467% | 2.773% | 2.232% | 1.576% |  | 1.263% | 4.41% | 2.155% | 3.472% | 7.356% | 3.786% | 7.088% | 2.884% | 3.13% | 1.096% | 4.224% | 0.807% | 2.098% | 1.843%
| PDM3 |  | 5% | 3.653% | -5.111% | 10.47% | 10.91% | 11.54% | 26.44% | 36.12% | 27.23% | 7.223% | 0.0313% | 0.2928% | -1.407% | -3.368% |  | -0.4917% | -9.514% | 36.88% | 16.11% | -13.44% | -1.93% | -16.64% | 8.586% | -31.35% | -2.192% | -14.74% | -4.415% | 4.216% | -26.26%
| Z-Score (Absolute) |  | 0.35 | 0.26 | 0.36 | 0.74 | 0.77 | 0.81 | 1.87 | 2.55 | 1.92 | 0.51 | 0 | 0.02 | 0.1 | 0.24 | 6.79 | 0.04 | 0.67 | 2.6 | 1.14 | 0.95 | 0.14 | 1.18 | 0.6 | 2.22 | 0.16 | 1.04 | 0.31 | 0.3 | 1.86
| Analytical results for W in OREAS 234b (Certified Value 23.5 ppm) | [32 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 27.2 | 21.19 | 28.63 | 27.7203 | 19.8 | 26 | 15.8 | 18.11 | 20 | 24.6 | 23.5 | 25 | 24.6 | < 100 | 24 | 19 | 25.1 | 16 | 25 | 21 | 25.5 | 28.3 | 25.2 | 21.43 | 28.1 | [6 empty] | 1
| 1 | 2 | 26.4 | 21.05 | 30.32 | 28.7119 | 18.3 | 25 | 17.4 | 17.23 | 20 | 24.8 | 24.6 | 25.6 | 25.5 | < 100 | 24.1 | 18 | 25.7 | 18 | 25 | 21 | 25.5 | 28.7 | 26 | 21.13 | 29 | [6 empty] | 31
| 1 | 3 | 26.4 | 20.85 | 29.1 | 24.7275 | 19.5 | 27 | 15.1 | 17.8 | 19 | 24.6 | 24.4 | 26.3 | 25.2 | < 100 | 24.4 | 18 | 26.8 | 16.5 | 25 | 20 | 26.6 | 27.8 | 26.2 | 21.34 | 27.9 | [6 empty] | 16
| 1 | 4 | 27.1 | 20.82 | 29.33 | 28.0153 | 17.8 | 28 | 16.7 | 18.65 | 20 | 24.7 | 24.6 | 25.2 | 24.4 | < 100 | 24.1 | 18 | 25.2 | 17 | 25 | 20 | 25.4 | 28.4 | 25.8 | 20.44 | 28 | [6 empty] | 23.5262
| 1 | 5 | 26.6 | 21.01 | 28.43 | 29.0181 | 19.3 | 32 | 16.3 | 17.47 | 20 | 24.3 | 24.4 | 26.1 | 24 | < 100 | 24.4 | 18 | 26.8 | 16.6 | 25 | 20 | 25.3 | 27.7 | 26.1 | 21.64 | 28.4 | [6 empty] | 127
| 1 | 6 | 26.9 | 21.16 | 28.5 | 29.5498 | 19.1 | 29 | 16.7 | 16.65 | 20 | 24 | 26.5 | 24.3 | 24.7 | < 100 | 23.5 | 18 | 24.2 | 17.3 | 25 | 21 | 26.2 | 28.5 | 26.2 | 21.83 | 27.6
| Mean |  | 26.7667 | 21.0133 | 29.0517 | 27.9572 | 18.9667 | 27.8333 | 16.3333 | 17.6517 | 19.8333 | 24.5 | 24.6667 | 25.4167 | 24.7333 |  | 24.0833 | 18.1667 | 25.6333 | 16.9 | 25 | 20.5 | 25.75 | 28.2333 | 25.9167 | 21.3017 | 28.1667
| Median |  | 26.75 | 21.03 | 28.865 | 28.3636 | 19.2 | 27.5 | 16.5 | 17.635 | 20 | 24.6 | 24.5 | 25.4 | 24.65 |  | 24.1 | 18 | 25.45 | 16.8 | 25 | 20.5 | 25.5 | 28.35 | 26.05 | 21.385 | 28.05
| Std Dev. |  | 0.35024 | 0.15371 | 0.71513 | 1.7159 | 0.76333 | 2.4833 | 0.80166 | 0.69898 | 0.40825 | 0.29665 | 0.98725 | 0.7414 | 0.54283 |  | 0.33116 | 0.40825 | 1.025 | 0.69857 | 0 | 0.54772 | 0.5244 | 0.39833 | 0.38166 | 0.48651 | 0.48442
| Rel.Std.Dev. |  | 1.308% | 0.7315% | 2.462% | 6.138% | 4.025% | 8.922% | 4.908% | 3.96% | 2.058% | 1.211% | 4.002% | 2.917% | 2.195% |  | 1.375% | 2.247% | 3.999% | 4.134% | 0% | 2.672% | 2.037% | 1.411% | 1.473% | 2.284% | 1.72%
| PDM3 |  | 13.77% | -10.68% | 23.49% | 18.83% | -19.38% | 18.31% | -30.57% | -24.97% | -15.7% | 4.139% | 4.848% | 8.036% | 5.131% |  | 2.368% | -22.78% | 8.957% | -28.17% | 6.265% | -12.86% | 9.453% | 20.01% | 10.16% | -9.455% | 19.72%
| Z-Score (Absolute) |  | 0.4 | 0.73 | 0.85 | 0.63 | 1.13 | 0.61 | 1.65 | 1.39 | 0.96 | 0.05 | 0.01 | 0.13 | 0 | 4.96 | 0.13 | 1.29 | 0.18 | 1.54 | 0.05 | 0.83 | 0.2 | 0.69 | 0.23 | 0.67 | 0.67
| Analytical results for Y in OREAS 234b (Certified Value 10.6 ppm) | [32 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 10.48 | 11.04 | 10.09 | 10.5232 | 10.7644 | 11.5 | 11.2 | 13.8 | 10.54 | 10 | 10.6 | 10.55 | 11.4 | 9.59 | 10.8 | 9 | 11.13 | 10.2 | 8.9 | 11.35 | 9.14 | 9.9 | 11.19 | 9.03 | [7 empty] | 1
| 1 | 2 | 10.44 | 11.13 | 10.14 | 10.6989 | 10.7241 | 10.8 | 10.6 | 13.5 | 10.42 | 10 | 10.9 | 10.4 | 11.3 | 10.35 | 10.35 | 9 | 11.53 | 10.2 | 8.5 | 11.35 | 9.43 | 10.1 | 11.09 | 8.96 | [7 empty] | 32
| 1 | 3 | 10.4 | 10.99 | 10.57 | 10.5587 | 10.7614 | 11 | 11.1 | 13.6 | 10.38 | 9 | 10.8 | 10.55 | 11.2 | 11.05 | 10.75 | 9 | 11.15 | 10.4 | 8.5 | 11.2 | 9.29 | 10 | 11.22 | 9.02 | [7 empty] | 16
| 1 | 4 | 10.88 | 11.06 | 10.2 | 10.4531 | 10.7803 | 10.8 | 11 | 13.5 | 10.83 | 9 | 11 | 10.45 | 10.95 | 10.55 | 10.65 | 9 | 11.45 | 10.5 | 8.9 | 11.35 | 9.26 | 9.8 | 10.88 | 9.19 | [7 empty] | 10.6429
| 1 | 5 | 10.89 | 11.13 | 10.56 | 10.5512 | 10.8824 | 11 | 10.1 | 13.6 | 10.67 | 9 | 10.8 | 10.4 | 10.9 | 9.92 | 10.7 | 9 | 11.1 | 10.4 | 8.4 | 11.15 | 9.12 | 10 | 11.47 | 9.11 | [7 empty] | 128
| 1 | 6 | 10.79 | 11.1 | 10.09 | 10.5805 | 10.9777 | 11.3 | 10.5 | 13.7 | 10.31 | 9 | 10.7 | 10.05 | 10.7 | 10.1 | 10.45 | 10 | 11.64 | 10.4 | 8 | 11.1 | 9.19 | 10 | 11 | 8.89
| Mean |  | 10.6467 | 11.075 | 10.275 | 10.5609 | 10.8151 | 11.0667 | 10.75 | 13.6167 | 10.525 | 9.3333 | 10.8 | 10.4 | 11.075 | 10.26 | 10.6167 | 9.1667 | 11.3333 | 10.35 | 8.5333 | 11.25 | 9.2383 | 9.9667 | 11.1417 | 9.0333
| Median |  | 10.635 | 11.08 | 10.17 | 10.555 | 10.7724 | 11 | 10.8 | 13.6 | 10.48 | 9 | 10.8 | 10.425 | 11.075 | 10.225 | 10.675 | 9 | 11.3 | 10.4 | 8.5 | 11.275 | 9.225 | 10 | 11.14 | 9.025
| Std Dev. |  | 0.23045 | 0.055408 | 0.22828 | 0.080717 | 0.095801 | 0.28048 | 0.42308 | 0.1169 | 0.19624 | 0.5164 | 0.14142 | 0.18439 | 0.26786 | 0.51116 | 0.17795 | 0.40825 | 0.23483 | 0.12247 | 0.33862 | 0.11402 | 0.11479 | 0.10328 | 0.20371 | 0.10633
| Rel.Std.Dev. |  | 2.165% | 0.5003% | 2.222% | 0.7643% | 0.8858% | 2.534% | 3.936% | 0.8585% | 1.865% | 5.533% | 1.309% | 1.773% | 2.419% | 4.982% | 1.676% | 4.454% | 2.072% | 1.183% | 3.968% | 1.013% | 1.243% | 1.036% | 1.828% | 1.177%
| PDM3 |  | 0.03496% | 4.06% | -3.457% | -0.7707% | 1.617% | 3.981% | 1.006% | 27.94% | -1.108% | -12.3% | 1.476% | -2.283% | 4.06% | -3.598% | -0.2469% | -13.87% | 6.487% | -2.752% | -19.82% | 5.704% | -13.2% | -6.354% | 4.686% | -15.12%
| Z-Score (Absolute) |  | 0.03 | 0.74 | 0.6 | 0.12 | 0.31 | 0.73 | 0.2 | 4.99 | 0.18 | NA | 0.28 | 0.39 | 0.74 | 0.62 | 0.03 | NA | 1.17 | 0.47 | 3.51 | 1.03 | 2.33 | 1.11 | 0.85 | 2.67
Results from laboratories 1.12 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 234b (Certified Value 0.97 ppm) | [32 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 2
| 1 | 1 | 1 | 1.026 | 0.811 | 0.95206 | 1.1 | 1.26 | 0.85 | 1 | [23 empty] | 1
| 1 | 2 | 1 | 1.076 | 0.814 | 0.93794 | 1.1 | 1.19 | 0.75 | 0.9 | [23 empty] | 33
| 1 | 3 | 1 | 1.034 | 0.916 | 0.92707 | 1.1 | 1.17 | 0.75 | 0.9 | [23 empty] | 16
| 1 | 4 | 1 | 1.037 | 0.866 | 0.90063 | 1.1 | 1.24 | 0.8 | 1 | [23 empty] | 0.97326
| 1 | 5 | 1 | 1.055 | 0.867 | 0.91922 | 1.1 | 1.1 | 0.75 | 0.9 | [23 empty] | 129
| 1 | 6 | 1 | 1.044 | 0.885 | 0.90881 | 1.1 | 1.14 | 0.7 | 0.9
| Mean |  | 1 | 1.0453 | 0.85983 | 0.92429 | 1.1 | 1.1833 | 0.76667 | 0.93333
| Median |  | 1 | 1.0405 | 0.8665 | 0.92314 | 1.1 | 1.18 | 0.75 | 0.9
| Std Dev. |  | 0 | 0.017929 | 0.040897 | 0.018925 | 2.4324e-16 | 0.060222 | 0.05164 | 0.05164
| Rel.Std.Dev. |  | 0% | 1.715% | 4.756% | 2.047% | 2.211e-14% | 5.089% | 6.736% | 5.533%
| PDM3 |  | 2.748% | 7.406% | -11.65% | -5.031% | 13.02% | 21.59% | -21.23% | -4.102%
| Z-Score (Absolute) |  | NA | 0.67 | 0.44 | 0.05 | 1 | 1.51 | 1 | 0
| Analytical results for Zn in OREAS 234b (Certified Value 129 ppm) | [32 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 132 | 128 | 112.69 | 124.2 | 133.4982 | 133 | 131 | 128 | 131 | 124 | 139 | 136 | 134 | 133 | < 100 | 147 | 120 | 129.256 | 132 | 116.7 | 129 | 125.26 | 130.25 | 125 | 139.4424 | 134 | 115.83 | 126 | 135 | 123 | 128 | 1
| 1 | 2 | 131 | 127.2 | 111.77 | 127.45 | 133.9396 | 127 | 126 | 134 | 127 | 123 | 144 | 137 | 131 | 135 | < 100 | 142 | 119 | 125.0248 | 134 | 127.4 | 131 | 127.46 | 130.18 | 120 | 140.3137 | 134 | 115.214 | 126 | 139 | 122 | 130 | 34
| 1 | 3 | 132 | 127.1 | 114.84 | 126.46 | 135.621 | 131 | 132 | 134 | 130 | 122 | 140 | 136 | 131 | 134 | < 100 | 144 | 120 | 128.4766 | 135 | 125.2 | 128 | 123.45 | 124.74 | 125 | 141.5648 | 135 | 123.344 | 121 | 135 | 123 | 128 | 16
| 1 | 4 | 133 | 127.4 | 113.25 | 123.5 | 131.9361 | 127 | 130 | 132 | 131 | 123 | 140 | 135 | 132 | 134 | < 100 | 145 | 117 | 127.2636 | 139 | 125.7 | 125 | 127.86 | 126.99 | 125 | 143.3305 | 135 | 123.633 | 129 | 130 | 122 | 130 | 129.4807
| 1 | 5 | 132 | 128.1 | 114.79 | 123.82 | 135.3463 | 128 | 128 | 131 | 132 | 124 | 139 | 135 | 133 | 133 | < 100 | 144 | 116 | 129.7152 | 143 | 123.4 | 134 | 121.83 | 129.82 | 125 | 140.6818 | 134 | 118.635 | 123 | 137 | 125 | 131 | 130
| 1 | 6 | 140 | 127.7 | 112.23 | 125.55 | 135.9795 | 131 | 126 | 132 | 127 | 124 | 138 | 134 | 131 | 131 | < 100 | 143 | 123 | 127.0684 | 140 | 128.7 | 130 | 124.24 | 129.89 | 120 | 137.6684 | 134 | 132.365 | 123 | 132 | 122 | 125
| Mean |  | 133.3333 | 127.5833 | 113.2617 | 125.1633 | 134.3868 | 129.5 | 128.8333 | 131.8333 | 129.6667 | 123.3333 | 140 | 135.5 | 132 | 133.3333 |  | 144.1667 | 119.1667 | 127.8008 | 137.1667 | 124.5167 | 129.5 | 125.0167 | 128.645 | 123.3333 | 140.5003 | 134.3333 | 121.5035 | 124.6667 | 134.6667 | 122.8333 | 128.6667
| Median |  | 132 | 127.55 | 112.97 | 124.875 | 134.6429 | 129.5 | 129 | 132 | 130.5 | 123.5 | 139.5 | 135.5 | 131.5 | 133.5 |  | 144 | 119.5 | 127.8701 | 137 | 125.45 | 129.5 | 124.75 | 129.855 | 125 | 140.4977 | 134 | 120.9895 | 124.5 | 135 | 122.5 | 129
| Std Dev. |  | 3.3267 | 0.41673 | 1.2995 | 1.5849 | 1.5477 | 2.51 | 2.5626 | 2.2286 | 2.1602 | 0.8165 | 2.0976 | 1.0488 | 1.2649 | 1.3663 |  | 1.7224 | 2.4833 | 1.7187 | 4.1673 | 4.2433 | 3.0166 | 2.3378 | 2.2738 | 2.582 | 1.9157 | 0.5164 | 6.4167 | 2.8752 | 3.266 | 1.169 | 2.1602
| Rel.Std.Dev. |  | 2.495% | 0.3266% | 1.147% | 1.266% | 1.152% | 1.938% | 1.989% | 1.69% | 1.666% | 0.662% | 1.498% | 0.774% | 0.9583% | 1.025% |  | 1.195% | 2.084% | 1.345% | 3.038% | 3.408% | 2.329% | 1.87% | 1.768% | 2.094% | 1.363% | 0.3844% | 5.281% | 2.306% | 2.425% | 0.9517% | 1.679%
| PDM3 |  | 2.975% | -1.465% | -12.53% | -3.334% | 3.789% | 0.01487% | -0.5% | 1.817% | 0.1436% | -4.748% | 8.124% | 4.649% | 1.946% | 2.975% |  | 11.34% | -7.966% | -1.297% | 5.936% | -3.834% | 0.01487% | -3.448% | -0.6455% | -4.748% | 8.511% | 3.748% | -6.161% | -3.718% | 4.005% | -5.134% | -0.6287%
| Z-Score (Absolute) |  | 0.67 | 0.19 | 2.33 | 0.55 | 0.83 | 0.1 | 0 | 0.45 | 0.12 | 0.82 | 1.67 | 1 | 0.47 | 0.67 | 11.81 | 2.3 | 1.45 | 0.15 | 1.25 | 0.65 | 0.1 | 0.57 | 0.03 | 0.82 | 1.75 | 0.82 | 1.1 | 0.62 | 0.87 | 0.9 | 0.02
| Analytical results for Zr in OREAS 234b (Certified Value 13.9 ppm) | [32 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 12.3 | 14 | 18.65 | 14.4914 | 13.4 | 12.2 | 16.6 | 18.2 | 13 | 10.6 | 13 | 15 | 12.4 | 13.5 | 12 | 13.6 | 11.5 | 17.028 | 8 | 15.2 | 10.7 | 12.7 | 14.7 | 10.7 | [7 empty] | 1
| 1 | 2 | 12.1 | 14.31 | 18.8 | 14.674 | 11.3 | 9.84 | 16.6 | 17.3 | 13 | 10.9 | 12.9 | 15.2 | 13.6 | 13.7 | 13 | 14.5 | 11.8 | 16.308 | 8 | 14.8 | 11.5 | 13.1 | 15 | 9.9 | [7 empty] | 35
| 1 | 3 | 12 | 13.96 | 19.6 | 14.7157 | 12.1 | 12.5 | 16.2 | 17.5 | 12 | 11.1 | 13.5 | 15.8 | 14.2 | 14.3 | 13 | 13.7 | 12.4 | 15.66 | 7 | 14.5 | 11.9 | 13 | 14.9 | 10.2 | [7 empty] | 16
| 1 | 4 | 12 | 14.04 | 20.62 | 14.5847 | 11.1 | 12.2 | 16.3 | 18 | 12 | 11.7 | 13.3 | 16.6 | 13.2 | 13.9 | 11 | 13.9 | 12.5 | 16.254 | 7 | 14.7 | 11.7 | 12.6 | 14.2 | 10.4 | [7 empty] | 13.9286
| 1 | 5 | 12.5 | 14.22 | 19.87 | 14.8917 | 12.3 | 12.5 | 15.7 | 17.8 | 12 | 10.7 | 13.1 | 15.9 | 12.6 | 14.2 | 12 | 14.4 | 12.4 | 17.343 | 7 | 14.5 | 11 | 13 | 15.2 | 10.6 | [7 empty] | 131
| 1 | 6 | 12.2 | 14.46 | 18.82 | 14.7864 | 11.4 | 12.6 | 15.6 | 17.1 | 12 | 11 | 13 | 15.2 | 13.2 | 13.6 | 13 | 14.7 | 12.4 | 17.316 | 7 | 14.7 | 10.9 | 13.2 | 14.5 | 10.1
| Mean |  | 12.1833 | 14.165 | 19.3933 | 14.6906 | 11.9333 | 11.9733 | 16.1667 | 17.65 | 12.3333 | 11 | 13.1333 | 15.6167 | 13.2 | 13.8667 | 12.3333 | 14.1333 | 12.1667 | 16.6515 | 7.3333 | 14.7333 | 11.2833 | 12.9333 | 14.75 | 10.3167
| Median |  | 12.15 | 14.13 | 19.21 | 14.6948 | 11.75 | 12.35 | 16.25 | 17.65 | 12 | 10.95 | 13.05 | 15.5 | 13.2 | 13.8 | 12.5 | 14.15 | 12.4 | 16.668 | 7 | 14.7 | 11.25 | 13 | 14.8 | 10.3
| Std Dev. |  | 0.19408 | 0.19796 | 0.77562 | 0.14241 | 0.85946 | 1.0584 | 0.43205 | 0.42308 | 0.5164 | 0.38987 | 0.22509 | 0.60139 | 0.65727 | 0.3266 | 0.8165 | 0.45898 | 0.41312 | 0.68126 | 0.5164 | 0.2582 | 0.48339 | 0.23381 | 0.36194 | 0.30605
| Rel.Std.Dev. |  | 1.593% | 1.398% | 3.999% | 0.9694% | 7.202% | 8.84% | 2.672% | 2.397% | 4.187% | 3.544% | 1.714% | 3.851% | 4.979% | 2.355% | 6.62% | 3.248% | 3.395% | 4.091% | 7.042% | 1.752% | 4.284% | 1.808% | 2.454% | 2.967%
| PDM3 |  | -12.53% | 1.697% | 39.23% | 5.471% | -14.32% | -14.04% | 16.07% | 26.72% | -11.45% | -21.03% | -5.71% | 12.12% | -5.231% | -0.4446% | -11.45% | 1.47% | -12.65% | 19.55% | -47.35% | 5.778% | -18.99% | -7.145% | 5.897% | -25.93%
| Z-Score (Absolute) |  | 0.67 | 0.12 | 2.19 | 0.33 | 0.77 | 0.75 | 0.91 | 1.5 | NA | 1.14 | 0.29 | 0.69 | 0.26 | 0 | NA | 0.11 | 0.67 | 1.1 | NA | 0.34 | 1.02 | 0.37 | 0.35 | 1.41
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C965: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1149: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1204: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 234b (Indicative Value 12.81 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.81 | 1
| 1 | 2 | 12.81 | 10
| Mean |  | 12.81 | 16
| Median |  | 12.81 | 12.81
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 234b (Indicative Value 8.94 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 8.93 | 1
| 1 | 2 | 8.95 | 11
| Mean |  | 8.94 | 16
| Median |  | 8.94 | 8.94
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.1582%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 234b (Indicative Value 11.62 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.61 | 1
| 1 | 2 | 11.62 | 12
| Mean |  | 11.615 | 16
| Median |  | 11.615 | 11.615
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.06088%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 234b (Indicative Value 0.673 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.671 | 1
| 1 | 2 | 0.675 | 13
| Mean |  | 0.673 | 16
| Median |  | 0.673 | 0.673
| Std Dev. |  | 0.0028284 | 19
| Rel.Std.Dev. |  | 0.4203%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 234b (Indicative Value 6.43 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.43 | 1
| 1 | 2 | 6.43 | 10
| Mean |  | 6.43 | 16
| Median |  | 6.43 | 6.43
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 234b (Indicative Value 0.18 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.181 | 1
| 1 | 2 | 0.179 | 11
| Mean |  | 0.18 | 16
| Median |  | 0.18 | 0.18
| Std Dev. |  | 0.0014142 | 17
| Rel.Std.Dev. |  | 0.7857%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 234b (Indicative Value 2.61 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.61 | 1
| 1 | 2 | 2.6 | 12
| Mean |  | 2.605 | 16
| Median |  | 2.605 | 2.605
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.2714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 234b (Indicative Value 0.095 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.095 | 1
| 1 | 2 | 0.095 | 13
| Mean |  | 0.095 | 16
| Median |  | 0.095 | 0.095
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 234b (Indicative Value 0.333 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.3308 | 1
| 1 | 2 | 0.336 | 17
| Mean |  | 0.3334 | 16
| Median |  | 0.3334 | 0.33337
| Std Dev. |  | 0.003677 | 16
| Rel.Std.Dev. |  | 1.103%
| PDM3 |  | 0.007612%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 234b (Indicative Value 52.19 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 52.19 | 1
| 1 | 2 | 52.19 | 11
| Mean |  | 52.19 | 16
| Median |  | 52.19 | 52.19
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 234b (Indicative Value 1.04 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.033 | 1
| 1 | 2 | 1.04 | 12
| Mean |  | 1.0365 | 16
| Median |  | 1.0365 | 1.0365
| Std Dev. |  | 0.0049497 | 18
| Rel.Std.Dev. |  | 0.4775%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 234b (Indicative Value 3.13 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.12 | 1
| 1 | 2 | 3.14 | 15
| Mean |  | 3.13 | 16
| Median |  | 3.13 | 3.13
| Std Dev. |  | 0.014142 | 21
| Rel.Std.Dev. |  | 0.4518%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 234b (Indicative Value 0.14 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.14 | 1
| 1 | 2 | 0.14 | 17
| Mean |  | 0.14 | 16
| Median |  | 0.14 | 0.14
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 234b (Indicative Value 0.375 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.38 | 17
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 234b (Indicative Value 0.35 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.3 | 19
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 234b (Indicative Value 63 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 62.6 | 1
| 1 | 2 | 64 | 20
| Mean |  | 63.3 | 16
| Median |  | 63.3 | 63.3
| Std Dev. |  | 0.98995 | 26
| Rel.Std.Dev. |  | 1.564%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 234b (Indicative Value 287 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 282 | 1
| 1 | 2 | 291 | 21
| Mean |  | 286.5 | 16
| Median |  | 286.5 | 286.5
| Std Dev. |  | 6.364 | 27
| Rel.Std.Dev. |  | 2.221%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 234b (Indicative Value 0.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 22
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 234b (Indicative Value 0.06 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.06 | 23
| Mean |  | 0.06 | 16
| Median |  | 0.06 | 0.06
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 234b (Indicative Value 0.7 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.7 | 24
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 234b (Indicative Value 12.6 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.7 | 1
| 1 | 2 | 12.4 | 25
| Mean |  | 12.55 | 16
| Median |  | 12.55 | 12.55
| Std Dev. |  | 0.21213 | 31
| Rel.Std.Dev. |  | 1.69%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 234b (Indicative Value 43.9 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 44 | 1
| 1 | 2 | 43.8 | 26
| Mean |  | 43.9 | 16
| Median |  | 43.9 | 43.9
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 0.3221%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 234b (Indicative Value 167 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 166 | 1
| 1 | 2 | 168 | 27
| Mean |  | 167 | 16
| Median |  | 167 | 167
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 0.8468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 234b (Indicative Value 0.91 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.88 | 1
| 1 | 2 | 0.94 | 28
| Mean |  | 0.91 | 16
| Median |  | 0.91 | 0.91
| Std Dev. |  | 0.042426 | 34
| Rel.Std.Dev. |  | 4.662%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 234b (Indicative Value 167 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 168 | 1
| 1 | 2 | 166 | 29
| Mean |  | 167 | 16
| Median |  | 167 | 167
| Std Dev. |  | 1.4142 | 35
| Rel.Std.Dev. |  | 0.8468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 234b (Indicative Value 3.6 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.69 | 1
| 1 | 2 | 3.51 | 30
| Mean |  | 3.6 | 16
| Median |  | 3.6 | 3.6
| Std Dev. |  | 0.12728 | 36
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 234b (Indicative Value 2.37 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.34 | 1
| 1 | 2 | 2.4 | 31
| Mean |  | 2.37 | 16
| Median |  | 2.37 | 2.37
| Std Dev. |  | 0.042426 | 37
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 234b (Indicative Value 0.89 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.88 | 1
| 1 | 2 | 0.89 | 32
| Mean |  | 0.885 | 16
| Median |  | 0.885 | 0.885
| Std Dev. |  | 0.0070711 | 38
| Rel.Std.Dev. |  | 0.799%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 234b (Indicative Value 14.7 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.7 | 1
| 1 | 2 | 14.6 | 33
| Mean |  | 14.65 | 16
| Median |  | 14.65 | 14.65
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.4827%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 234b (Indicative Value 3.12 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.1 | 1
| 1 | 2 | 3.14 | 34
| Mean |  | 3.12 | 16
| Median |  | 3.12 | 3.12
| Std Dev. |  | 0.028284 | 40
| Rel.Std.Dev. |  | 0.9065%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 234b (Indicative Value 1.43 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.45 | 35
| Mean |  | 1.425 | 16
| Median |  | 1.425 | 1.425
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 234b (Indicative Value 1.87 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.83 | 1
| 1 | 2 | 1.91 | 19
| Mean |  | 1.87 | 16
| Median |  | 1.87 | 1.87
| Std Dev. |  | 0.056569 | 25
| Rel.Std.Dev. |  | 3.025%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 234b (Indicative Value 0.8 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.81 | 1
| 1 | 2 | 0.79 | 20
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 1.768%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 234b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 21
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 234b (Indicative Value 5.63 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.66 | 1
| 1 | 2 | 5.6 | 22
| Mean |  | 5.63 | 16
| Median |  | 5.63 | 5.63
| Std Dev. |  | 0.042426 | 28
| Rel.Std.Dev. |  | 0.7536%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 234b (Indicative Value 0.33 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.33 | 23
| Mean |  | 0.325 | 16
| Median |  | 0.325 | 0.325
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 2.176%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 234b (Indicative Value 0.157 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.156 | 1
| 1 | 2 | 0.157 | 24
| Mean |  | 0.1565 | 16
| Median |  | 0.1565 | 0.1565
| Std Dev. |  | 0.00070711 | 30
| Rel.Std.Dev. |  | 0.4518%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 234b (Indicative Value 2.4 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.4 | 25
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 234b (Indicative Value 3.62 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.59 | 1
| 1 | 2 | 3.64 | 8
| Mean |  | 3.615 | 16
| Median |  | 3.615 | 3.615
| Std Dev. |  | 0.035355 | 32
| Rel.Std.Dev. |  | 0.978%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 234b (Indicative Value 8.53 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.41 | 1
| 1 | 2 | 8.64 | 27
| Mean |  | 8.525 | 16
| Median |  | 8.525 | 8.525
| Std Dev. |  | 0.16263 | 33
| Rel.Std.Dev. |  | 1.908%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 234b (Indicative Value 102 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 102 | 1
| 1 | 2 | 102 | 28
| Mean |  | 102 | 16
| Median |  | 102 | 102
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 234b (Indicative Value 26.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 27 | 29
| Mean |  | 26.5 | 16
| Median |  | 26.5 | 26.5
| Std Dev. |  | 0.70711 | 35
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 234b (Indicative Value 1.85 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.86 | 1
| 1 | 2 | 1.83 | 30
| Mean |  | 1.845 | 16
| Median |  | 1.845 | 1.845
| Std Dev. |  | 0.021213 | 36
| Rel.Std.Dev. |  | 1.15%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 234b (Indicative Value 14.5 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.6 | 1
| 1 | 2 | 14.4 | 31
| Mean |  | 14.5 | 16
| Median |  | 14.5 | 14.5
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 0.9753%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 234b (Indicative Value 0.013 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | < 0.01 | 3
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.0125
| Std Dev. |  |  | 38
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 234b (Indicative Value 1.5 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.6 | 33
| Mean |  | 1.5 | 16
| Median |  | 1.5 | 1.5
| Std Dev. |  | 0.14142 | 39
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 234b (Indicative Value 38 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.3 | 1
| 1 | 2 | 37.7 | 34
| Mean |  | 38 | 16
| Median |  | 38 | 38
| Std Dev. |  | 0.42426 | 40
| Rel.Std.Dev. |  | 1.116%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 234b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 4
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 234b (Indicative Value 2.48 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.49 | 1
| 1 | 2 | 2.46 | 19
| Mean |  | 2.475 | 16
| Median |  | 2.475 | 2.475
| Std Dev. |  | 0.021213 | 25
| Rel.Std.Dev. |  | 0.8571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 234b (Indicative Value 1 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1 | 20
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 234b (Indicative Value 188 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 189 | 1
| 1 | 2 | 187 | 21
| Mean |  | 188 | 16
| Median |  | 188 | 188
| Std Dev. |  | 1.4142 | 27
| Rel.Std.Dev. |  | 0.7522%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 234b (Indicative Value 0.24 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.25 | 22
| Mean |  | 0.235 | 16
| Median |  | 0.235 | 0.235
| Std Dev. |  | 0.021213 | 28
| Rel.Std.Dev. |  | 9.027%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 234b (Indicative Value 0.57 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.57 | 1
| 1 | 2 | 0.56 | 23
| Mean |  | 0.565 | 16
| Median |  | 0.565 | 0.565
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 234b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 24
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 234b (Indicative Value 1.01 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.04 | 1
| 1 | 2 | 0.98 | 25
| Mean |  | 1.01 | 16
| Median |  | 1.01 | 1.01
| Std Dev. |  | 0.042426 | 31
| Rel.Std.Dev. |  | 4.201%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 234b (Indicative Value 0.618 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.614 | 1
| 1 | 2 | 0.622 | 26
| Mean |  | 0.618 | 16
| Median |  | 0.618 | 0.618
| Std Dev. |  | 0.0056569 | 32
| Rel.Std.Dev. |  | 0.9153%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 234b (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 234b (Indicative Value 0.35 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.34 | 28
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 234b (Indicative Value 0.31 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.31 | 1
| 1 | 2 | 0.31 | 29
| Mean |  | 0.31 | 16
| Median |  | 0.31 | 0.31
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 234b (Indicative Value 268 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 268 | 1
| 1 | 2 | 268 | 30
| Mean |  | 268 | 16
| Median |  | 268 | 268
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 234b (Indicative Value 32.8 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33 | 1
| 1 | 2 | 32.5 | 31
| Mean |  | 32.75 | 16
| Median |  | 32.75 | 32.75
| Std Dev. |  | 0.35355 | 37
| Rel.Std.Dev. |  | 1.08%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 234b (Indicative Value 20.1 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.1 | 1
| 1 | 2 | 20.1 | 32
| Mean |  | 20.1 | 16
| Median |  | 20.1 | 20.1
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 234b (Indicative Value 2.25 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.19 | 1
| 1 | 2 | 2.31 | 33
| Mean |  | 2.25 | 16
| Median |  | 2.25 | 2.25
| Std Dev. |  | 0.084853 | 39
| Rel.Std.Dev. |  | 3.771%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 234b (Indicative Value 173 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 145 | 1
| 1 | 2 | 200 | 34
| Mean |  | 172.5 | 16
| Median |  | 172.5 | 172.5
| Std Dev. |  | 38.8909 | 40
| Rel.Std.Dev. |  | 22.55%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 234b (Indicative Value 65 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 65 | 1
| 1 | 2 | 65.5 | 35
| Mean |  | 65.25 | 16
| Median |  | 65.25 | 65.25
| Std Dev. |  | 0.35355 | 41
| Rel.Std.Dev. |  | 0.5418%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.