## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 240b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 5.6479 | 5.5986 | 5.6972 | 5.6221 | 5.6737
PhotonAssay (recommended gross mass* 315±15 g)
Au, Gold (ppm) | 5.575 | 5.5057 | 5.6444 | 5.5681 | 5.582
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 5.5261 | 5.4521 | 5.6001 | 5.4985 | 5.5537
Cyanide Leach
Au, Gold (ppm) | 5.4353 | 5.3661 | 5.5046 | 5.4257 | 5.4449
4-Acid Digestion
Ag, Silver (ppm) | 1.3996 | 1.3403 | 1.4588 | 1.3388 | 1.4603
Al, Aluminium (wt.%) | 6.4931 | 6.3082 | 6.6779 | 6.337 | 6.6491
As, Arsenic (ppm) | 97.988 | 94.4973 | 101.4786 | 94.2082 | 101.7678
Ba, Barium (ppm) | 323.8994 | 314.9844 | 332.8144 | 313.0343 | 334.7646
Be, Beryllium (ppm) | 0.47001 | 0.42556 | 0.51446 | 0.45081 | 0.48921
Bi, Bismuth (ppm) | 0.072995 | 0.060331 | 0.085659 | IND | IND
Ca, Calcium (wt.%) | 5.6632 | 5.4991 | 5.8273 | 5.5655 | 5.7609
Cd, Cadmium (ppm) | 0.88591 | 0.82465 | 0.94717 | 0.84552 | 0.9263
Ce, Cerium (ppm) | 14.8792 | 14.1641 | 15.5944 | 14.3951 | 15.3634
Co, Cobalt (ppm) | 40.2925 | 39.0581 | 41.5269 | 39.0278 | 41.5572
Cr, Chromium (ppm) | 145.6833 | 137.5269 | 153.8396 | 141.0679 | 150.2987
Cs, Caesium (ppm) | 1.3785 | 1.31 | 1.4471 | 1.3393 | 1.4177
Cu, Copper (ppm) | 164.3747 | 159.1948 | 169.5546 | 160.3779 | 168.3714
Dy, Dysprosium (ppm) | 3.3548 | 3.0529 | 3.6567 | 3.2371 | 3.4724
Er, Erbium (ppm) | 2.0298 | 1.8283 | 2.2313 | 1.8863 | 2.1734
Eu, Europium (ppm) | 0.8694 | 0.79634 | 0.94246 | 0.82057 | 0.91823
Fe, Iron (wt.%) | 7.3127 | 7.124 | 7.5014 | 7.1664 | 7.4591
Ga, Gallium (ppm) | 14.8328 | 14.1216 | 15.544 | 14.274 | 15.3917
Gd, Gadolinium (ppm) | 2.8855 | 2.6005 | 3.1706 | 2.7345 | 3.0366
Ge, Germanium (ppm) | 0.0953 | 0.069222 | 0.12138 | IND | IND
Hf, Hafnium (ppm) | 1.6641 | 1.5338 | 1.7944 | 1.5796 | 1.7485
Ho, Holmium (ppm) | 0.67009 | 0.59052 | 0.74965 | 0.63897 | 0.70121
In, Indium (ppm) | 0.075469 | 0.068029 | 0.082909 | 0.067305 | 0.083633
K, Potassium (wt.%) | 0.72952 | 0.70667 | 0.75237 | 0.70562 | 0.75343
La, Lanthanum (ppm) | 6.6049 | 6.2873 | 6.9225 | 6.3997 | 6.8101
Li, Lithium (ppm) | 13.4658 | 12.9229 | 14.0087 | 13.1121 | 13.8195
Lu, Lutetium (ppm) | 0.2946 | 0.2569 | 0.33231 | 0.27741 | 0.31179
Mg, Magnesium (wt.%) | 3.4378 | 3.3566 | 3.5189 | 3.3611 | 3.5144
Mn, Manganese (wt.%) | 0.1396 | 0.13559 | 0.1436 | 0.13707 | 0.14213
Mo, Molybdenum (ppm) | 2.9033 | 2.7521 | 3.0546 | 2.7724 | 3.0343
Na, Sodium (wt.%) | 1.8708 | 1.8284 | 1.9132 | 1.8126 | 1.929
Nb, Niobium (ppm) | 3.5405 | 3.316 | 3.7649 | 3.358 | 3.7229
Nd, Neodymium (ppm) | 8.742 | 8.1503 | 9.3337 | 8.403 | 9.081
Ni, Nickel (ppm) | 94.3142 | 91.0452 | 97.5832 | 92.2143 | 96.4141
P, Phosphorus (wt.%) | 0.042884 | 0.041671 | 0.044097 | 0.041924 | 0.043843
Pb, Lead (ppm) | 40.5454 | 38.5117 | 42.5791 | 39.0994 | 41.9913
Pr, Praseodymium (ppm) | 1.9111 | 1.7664 | 2.0558 | 1.8289 | 1.9933
Rb, Rubidium (ppm) | 21.059 | 20.1351 | 21.9829 | 20.3241 | 21.794
Re, Rhenium (ppm) | 0.0024963 | 0.001751 | 0.0032416 | IND | IND
S, Sulphur (wt.%) | 0.54945 | 0.53547 | 0.56344 | 0.53725 | 0.56166
Sb, Antimony (ppm) | 2.1393 | 2.006 | 2.2727 | 2.0048 | 2.2739
Sc, Scandium (ppm) | 33.328 | 32.1161 | 34.54 | 32.3012 | 34.3548
Sm, Samarium (ppm) | 2.4241 | 2.186 | 2.6621 | 2.2812 | 2.5669
Sn, Tin (ppm) | 1.3165 | 1.2021 | 1.4308 | 1.1143 | 1.5186
Sr, Strontium (ppm) | 196.3146 | 190.4191 | 202.21 | 191.8416 | 200.7876
Ta, Tantalum (ppm) | 0.24205 | 0.21624 | 0.26786 | 0.22439 | 0.25971
Tb, Terbium (ppm) | 0.50369 | 0.4617 | 0.54568 | 0.4728 | 0.53458
Te, Tellurium (ppm) | 0.15299 | 0.12418 | 0.18179 | IND | IND
Th, Thorium (ppm) | 1.4503 | 1.3433 | 1.5573 | 1.3904 | 1.5103
Ti, Titanium (wt.%) | 0.54344 | 0.52907 | 0.55781 | 0.52768 | 0.55919
Tl, Thallium (ppm) | 0.34229 | 0.32238 | 0.36219 | 0.3213 | 0.36327
Tm, Thulium (ppm) | 0.29292 | 0.26152 | 0.32432 | 0.27037 | 0.31547
U, Uranium (ppm) | 0.42353 | 0.38821 | 0.45885 | 0.40036 | 0.44669
V, Vanadium (ppm) | 231.5352 | 225.0255 | 238.0449 | 226.9107 | 236.1596
W, Tungsten (ppm) | 47.1217 | 44.9608 | 49.2827 | 45.1301 | 49.1133
Y, Yttrium (ppm) | 18.2682 | 17.5459 | 18.9905 | 17.7255 | 18.8109
Yb, Ytterbium (ppm) | 1.8663 | 1.7013 | 2.0313 | 1.7569 | 1.9756
Zn, Zinc (ppm) | 173.4693 | 168.2596 | 178.679 | 169.0286 | 177.91
Zr, Zirconium (ppm) | 53.1248 | 49.8057 | 56.4439 | 50.5556 | 55.6939
Aqua Regia Digestion
Ag, Silver (ppm) | 1.3898 | 1.3375 | 1.442 | 1.3477 | 1.4318
Al, Aluminium (wt.%) | 2.9899 | 2.8749 | 3.1049 | 2.9205 | 3.0594
As, Arsenic (ppm) | 95.5824 | 92.6348 | 98.53 | 93.5578 | 97.607
B, Boron (ppm) | 22.5532 | 17.2557 | 27.8506 | 20.6959 | 24.4104
Ba, Barium (ppm) | 50.6824 | 47.8994 | 53.4654 | 49.5183 | 51.8465
Be, Beryllium (ppm) | 0.24784 | 0.21634 | 0.27934 | 0.22864 | 0.26703
Bi, Bismuth (ppm) | 0.067386 | 0.058999 | 0.075772 | IND | IND
Ca, Calcium (wt.%) | 2.2395 | 2.1206 | 2.3584 | 2.1865 | 2.2925
Cd, Cadmium (ppm) | 0.85933 | 0.80932 | 0.90934 | 0.82783 | 0.89083
Ce, Cerium (ppm) | 11.6936 | 11.177 | 12.2103 | 11.3997 | 11.9876
Co, Cobalt (ppm) | 29.8154 | 28.6984 | 30.9324 | 29.0832 | 30.5475
Cr, Chromium (ppm) | 34.6791 | 33.0615 | 36.2966 | 33.63 | 35.7282
Cs, Caesium (ppm) | 0.99546 | 0.95809 | 1.0328 | 0.96432 | 1.0266
Cu, Copper (ppm) | 162.498 | 158.147 | 166.8491 | 158.9553 | 166.0408
Dy, Dysprosium (ppm) | 1.9331 | 1.6034 | 2.2628 | 1.8474 | 2.0188
Er, Erbium (ppm) | 1.1141 | 0.92855 | 1.2996 | 1.0583 | 1.1698
Eu, Europium (ppm) | 0.45028 | 0.3682 | 0.53236 | 0.42934 | 0.47122
Fe, Iron (wt.%) | 5.3988 | 5.2891 | 5.5084 | 5.2898 | 5.5077
Ga, Gallium (ppm) | 9.3294 | 8.844 | 9.8149 | 9.124 | 9.5348
Gd, Gadolinium (ppm) | 1.8069 | 1.5624 | 2.0515 | 1.7388 | 1.8751
Ge, Germanium (ppm) | 0.094167 | 0.067368 | 0.12097 | IND | IND
Hf, Hafnium (ppm) | 0.43377 | 0.38919 | 0.47836 | 0.40623 | 0.46131
Ho, Holmium (ppm) | 0.37267 | 0.33079 | 0.41454 | 0.36152 | 0.38381
In, Indium (ppm) | 0.038395 | 0.034477 | 0.042313 | 0.035366 | 0.041424
K, Potassium (wt.%) | 0.20941 | 0.20094 | 0.21788 | 0.20309 | 0.21573
La, Lanthanum (ppm) | 5.4853 | 5.2856 | 5.6849 | 5.3584 | 5.6122
Li, Lithium (ppm) | 10.5696 | 10.1209 | 11.0183 | 10.3366 | 10.8027
Lu, Lutetium (ppm) | 0.12961 | 0.10453 | 0.15468 | IND | IND
Mg, Magnesium (wt.%) | 1.696 | 1.6455 | 1.7465 | 1.6582 | 1.7338
Mn, Manganese (wt.%) | 0.078831 | 0.075974 | 0.081689 | 0.077356 | 0.080307
Mo, Molybdenum (ppm) | 2.7855 | 2.6727 | 2.8984 | 2.6663 | 2.9047
Na, Sodium (wt.%) | 0.3283 | 0.31507 | 0.34154 | 0.31637 | 0.34024
Nd, Neodymium (ppm) | 6.2576 | 5.5543 | 6.9608 | 6.0857 | 6.4294
Ni, Nickel (ppm) | 67.9061 | 65.5959 | 70.2162 | 66.5424 | 69.2697
P, Phosphorus (wt.%) | 0.041627 | 0.040468 | 0.042786 | 0.040803 | 0.042452
Pb, Lead (ppm) | 40.6848 | 39.2303 | 42.1393 | 39.6096 | 41.76
Pr, Praseodymium (ppm) | 1.392 | 1.2289 | 1.5551 | 1.3513 | 1.4327
Rb, Rubidium (ppm) | 8.9095 | 8.5238 | 9.2952 | 8.6676 | 9.1514
Re, Rhenium (ppm) | 0.0019375 | 0.0012365 | 0.0026385 | IND | IND
S, Sulphur (wt.%) | 0.54527 | 0.53301 | 0.55753 | 0.53364 | 0.55689
Sb, Antimony (ppm) | 1.2113 | 1.0118 | 1.4109 | 1.1338 | 1.2889
Sc, Scandium (ppm) | 5.2114 | 4.8322 | 5.5906 | 5.0604 | 5.3624
Sm, Samarium (ppm) | 1.5952 | 1.4065 | 1.7838 | 1.5416 | 1.6487
Sn, Tin (ppm) | 0.74978 | 0.6846 | 0.81497 | 0.68035 | 0.81922
Sr, Strontium (ppm) | 43.2605 | 40.2492 | 46.2718 | 41.4784 | 45.0426
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.29002 | 0.2703 | 0.30975 | 0.27588 | 0.30417
Te, Tellurium (ppm) | 0.13814 | 0.12275 | 0.15353 | IND | IND
Th, Thorium (ppm) | 1.1379 | 1.059 | 1.2169 | 1.0858 | 1.19
Ti, Titanium (wt.%) | 0.28216 | 0.25751 | 0.30682 | 0.27207 | 0.29225
Tl, Thallium (ppm) | 0.18027 | 0.16869 | 0.19185 | 0.13204 | 0.2285
Tm, Thulium (ppm) | 0.14058 | 0.10667 | 0.17448 | IND | IND
U, Uranium (ppm) | 0.28384 | 0.26854 | 0.29914 | 0.26642 | 0.30125
V, Vanadium (ppm) | 115.6936 | 108.7092 | 122.6781 | 112.8669 | 118.5204
W, Tungsten (ppm) | 33.7297 | 31.1829 | 36.2764 | 32.837 | 34.6224
Y, Yttrium (ppm) | 10.0103 | 9.6356 | 10.385 | 9.8279 | 10.1927
Yb, Ytterbium (ppm) | 0.90634 | 0.79318 | 1.0195 | 0.80027 | 1.0124
Zn, Zinc (ppm) | 158.0831 | 152.8317 | 163.3345 | 155.553 | 160.6131
Zr, Zirconium (ppm) | 14.4571 | 13.2572 | 15.657 | 13.8222 | 15.0921
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 240b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 14.1126 | Pt | ppb | 8.0945
4-Acid Digestion
B | ppm | 33.7917 | Hg | ppm | 0.098133 | Se | ppm | 1.049
Aqua Regia Digestion
Hg | ppm | 0.062154 | Pd | ppb | 12.5 | Si | wt.% | 0.088959
Ir | ppm | < 0.003 | Pt | ppb | 6.7994
Nb | ppm | 0.11255 | Se | ppm | 0.73208
Borate Fusion XRF
Al2O3 | wt.% | 12.32 | MgO | wt.% | 5.86 | S | wt.% | 0.45651
CaO | wt.% | 8.16 | MnO | wt.% | 0.1685 | SiO2 | wt.% | 54.61
Fe2O3 | wt.% | 10.825 | Na2O | wt.% | 2.5 | TiO2 | wt.% | 0.957
K2O | wt.% | 0.8525 | P2O5 | wt.% | 0.0945
Thermogravimetry
LOI1000 | wt.% | 3.355
Infrared Combustion
C | wt.% | 0.19 | S | wt.% | 0.515
Laser Ablation ICP-MS
Ag | ppm | 1.3 | Hf | ppm | 2.02 | Sm | ppm | 2.535
As | ppm | 93.7 | Ho | ppm | 0.775 | Sn | ppm | 1.2
Ba | ppm | 313.5 | In | ppm | 0.075 | Sr | ppm | 187.5
Be | ppm | 0.7 | La | ppm | 6.69 | Ta | ppm | 0.225
Bi | ppm | 0.07 | Lu | ppm | 0.32 | Tb | ppm | 0.54
Cd | ppm | 0.95 | Mn | wt.% | 0.148 | Te | ppm | < 0.2
Ce | ppm | 14.35 | Mo | ppm | 2.9 | Th | ppm | 1.46
Co | ppm | 41.2 | Nb | ppm | 3.54 | Ti | wt.% | 0.566
Cr | ppm | 164.5 | Nd | ppm | 9.03 | Tl | ppm | 0.15
Cs | ppm | 1.32 | Ni | ppm | 98 | Tm | ppm | 0.33
Cu | ppm | 169 | Pb | ppm | 40.5 | U | ppm | 0.385
Dy | ppm | 3.445 | Pr | ppm | 1.99 | V | ppm | 240.5
Er | ppm | 2.195 | Rb | ppm | 20 | W | ppm | 46.25
Eu | ppm | 0.875 | Re | ppm | 0.02 | Y | ppm | 19.25
Ga | ppm | 14.3 | Sb | ppm | 2.2 | Yb | ppm | 2.075
Gd | ppm | 3.105 | Sc | ppm | 34.1 | Zn | ppm | 172.5
Ge | ppm | 1.45 | Se | ppm | < 5 | Zr | ppm | 70
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 240b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 5.6479 | 0.14263 | 5.3626 | 5.9332 | 5.22 | 6.0758 | 2.525% | 5.051% | 7.576% | 5.3655 | 5.9303
PhotonAssay
Au, ppm | 5.575 | 0.15704 | 5.2609 | 5.8891 | 5.1039 | 6.0461 | 2.817% | 5.634% | 8.451% | 5.2963 | 5.8538
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 5.5261 | 0.17109 | 5.1839 | 5.8683 | 5.0128 | 6.0393 | 3.096% | 6.192% | 9.288% | 5.2498 | 5.8024
Cyanide Leach
Au, ppm | 5.4353 | 0.14312 | 5.1491 | 5.7216 | 5.006 | 5.8647 | 2.633% | 5.266% | 7.899% | 5.1636 | 5.7071
4-Acid Digestion
Ag, ppm | 1.3996 | 0.048538 | 1.3025 | 1.4966 | 1.254 | 1.5452 | 3.468% | 6.936% | 10.4% | 1.3296 | 1.4695
Al, wt.% | 6.4931 | 0.27262 | 5.9478 | 7.0383 | 5.6752 | 7.3109 | 4.199% | 8.397% | 12.6% | 6.1684 | 6.8177
As, ppm | 97.988 | 4.5376 | 88.9128 | 107.0631 | 84.3752 | 111.6007 | 4.631% | 9.262% | 13.89% | 93.0886 | 102.8874
Ba, ppm | 323.8994 | 9.9736 | 303.9522 | 343.8466 | 293.9786 | 353.8202 | 3.079% | 6.158% | 9.238% | 307.7044 | 340.0944
Be, ppm | 0.47001 | 0.047954 | 0.3741 | 0.56592 | 0.32615 | 0.61387 | 10.2% | 20.41% | 30.61% | 0.44651 | 0.49351
Bi, ppm | 0.072995 | 0.010168 | 0.05266 | 0.093331 | 0.042492 | 0.1035 | 13.93% | 27.86% | 41.79% | 0.069345 | 0.076645
Ca, wt.% | 5.6632 | 0.21972 | 5.2237 | 6.1026 | 5.004 | 6.3223 | 3.88% | 7.76% | 11.64% | 5.38 | 5.9463
Cd, ppm | 0.88591 | 0.050394 | 0.78512 | 0.9867 | 0.73473 | 1.0371 | 5.688% | 11.38% | 17.07% | 0.84161 | 0.9302
Ce, ppm | 14.8792 | 0.84246 | 13.1943 | 16.5641 | 12.3518 | 17.4066 | 5.662% | 11.32% | 16.99% | 14.1353 | 15.6232
Co, ppm | 40.2925 | 1.6818 | 36.9289 | 43.6562 | 35.247 | 45.338 | 4.174% | 8.348% | 12.52% | 38.2779 | 42.3071
Cr, ppm | 145.6833 | 15.0102 | 115.6629 | 175.7036 | 100.6527 | 190.7138 | 10.3% | 20.61% | 30.91% | 138.3991 | 152.9674
Cs, ppm | 1.3785 | 0.083083 | 1.2124 | 1.5447 | 1.1293 | 1.6278 | 6.027% | 12.05% | 18.08% | 1.3096 | 1.4475
Cu, ppm | 164.3747 | 5.6296 | 153.1156 | 175.6338 | 147.486 | 181.2633 | 3.425% | 6.85% | 10.27% | 156.156 | 172.5934
Dy, ppm | 3.3548 | 0.10915 | 3.1365 | 3.5731 | 3.0273 | 3.6822 | 3.254% | 6.507% | 9.761% | 3.187 | 3.5225
Er, ppm | 2.0298 | 0.12768 | 1.7745 | 2.2852 | 1.6468 | 2.4129 | 6.29% | 12.58% | 18.87% | 1.9283 | 2.1313
Eu, ppm | 0.8694 | 0.053691 | 0.76202 | 0.97678 | 0.70833 | 1.0305 | 6.176% | 12.35% | 18.53% | 0.82593 | 0.91287
Fe, wt.% | 7.3127 | 0.33497 | 6.6428 | 7.9827 | 6.3078 | 8.3177 | 4.581% | 9.161% | 13.74% | 6.9471 | 7.6784
Ga, ppm | 14.8328 | 0.83538 | 13.162 | 16.5036 | 12.3267 | 17.339 | 5.632% | 11.26% | 16.9% | 14.0912 | 15.5744
Gd, ppm | 2.8855 | 0.25475 | 2.376 | 3.395 | 2.1213 | 3.6498 | 8.829% | 17.66% | 26.49% | 2.7413 | 3.0298
Ge, ppm | 0.0953 | 0.031236 | 0.032829 | 0.15777 | 0.0015928 | 0.18901 | 32.78% | 65.55% | 98.33% | 0.090535 | 0.10007
Hf, ppm | 1.6641 | 0.15304 | 1.358 | 1.9702 | 1.205 | 2.1232 | 9.196% | 18.39% | 27.59% | 1.5809 | 1.7473
Ho, ppm | 0.67009 | 0.037766 | 0.59456 | 0.74562 | 0.55679 | 0.78338 | 5.636% | 11.27% | 16.91% | 0.63658 | 0.70359
In, ppm | 0.075469 | 0.0061204 | 0.063228 | 0.08771 | 0.057108 | 0.09383 | 8.11% | 16.22% | 24.33% | 0.071696 | 0.079242
K, wt.% | 0.72952 | 0.030249 | 0.66902 | 0.79002 | 0.63878 | 0.82027 | 4.146% | 8.293% | 12.44% | 0.69305 | 0.766
La, ppm | 6.6049 | 0.38727 | 5.8304 | 7.3795 | 5.4431 | 7.7667 | 5.863% | 11.73% | 17.59% | 6.2747 | 6.9352
Li, ppm | 13.4658 | 0.91703 | 11.6317 | 15.2998 | 10.7147 | 16.2169 | 6.81% | 13.62% | 20.43% | 12.7925 | 14.1391
Lu, ppm | 0.2946 | 0.024921 | 0.24476 | 0.34444 | 0.21984 | 0.36936 | 8.459% | 16.92% | 25.38% | 0.27987 | 0.30933
Mg, wt.% | 3.4378 | 0.1517 | 3.1344 | 3.7412 | 2.9827 | 3.8929 | 4.413% | 8.826% | 13.24% | 3.2659 | 3.6096
Mn, wt.% | 0.1396 | 0.0039339 | 0.13173 | 0.14747 | 0.1278 | 0.1514 | 2.818% | 5.636% | 8.454% | 0.13262 | 0.14658
Mo, ppm | 2.9033 | 0.25234 | 2.3986 | 3.408 | 2.1463 | 3.6604 | 8.692% | 17.38% | 26.07% | 2.7582 | 3.0485
Na, wt.% | 1.8708 | 0.061171 | 1.7485 | 1.9932 | 1.6873 | 2.0543 | 3.27% | 6.539% | 9.809% | 1.7773 | 1.9644
Nb, ppm | 3.5405 | 0.22586 | 3.0887 | 3.9922 | 2.8629 | 4.218 | 6.379% | 12.76% | 19.14% | 3.3634 | 3.7175
Nd, ppm | 8.742 | 0.36086 | 8.0203 | 9.4637 | 7.6594 | 9.8246 | 4.128% | 8.256% | 12.38% | 8.3049 | 9.1791
Ni, ppm | 94.3142 | 3.8459 | 86.6224 | 102.0059 | 82.7765 | 105.8518 | 4.078% | 8.155% | 12.23% | 89.5985 | 99.0299
P, wt.% | 0.042884 | 0.0010045 | 0.040875 | 0.044893 | 0.03987 | 0.045897 | 2.342% | 4.685% | 7.027% | 0.040739 | 0.045028
Pb, ppm | 40.5454 | 2.8494 | 34.8466 | 46.2441 | 31.9973 | 49.0935 | 7.028% | 14.06% | 21.08% | 38.5181 | 42.5726
Pr, ppm | 1.9111 | 0.092772 | 1.7256 | 2.0967 | 1.6328 | 2.1894 | 4.854% | 9.709% | 14.56% | 1.8156 | 2.0067
Rb, ppm | 21.059 | 1.0981 | 18.8627 | 23.2553 | 17.7646 | 24.3534 | 5.215% | 10.43% | 15.64% | 20.0061 | 22.112
Re, ppm | 0.0024963 | 0.00065681 | 0.0011827 | 0.0038099 | 0.00052586 | 0.0044667 | 26.31% | 52.62% | 78.93% | 0.0023715 | 0.0026211
S, wt.% | 0.54945 | 0.015172 | 0.51911 | 0.5798 | 0.50394 | 0.59497 | 2.761% | 5.522% | 8.284% | 0.52198 | 0.57693
Sb, ppm | 2.1393 | 0.1686 | 1.8021 | 2.4765 | 1.6335 | 2.6451 | 7.881% | 15.76% | 23.64% | 2.0324 | 2.2463
Sc, ppm | 33.328 | 1.4324 | 30.4632 | 36.1929 | 29.0308 | 37.6253 | 4.298% | 8.596% | 12.89% | 31.6616 | 34.9944
Sm, ppm | 2.4241 | 0.13311 | 2.1578 | 2.6903 | 2.0247 | 2.8234 | 5.491% | 10.98% | 16.47% | 2.3029 | 2.5453
Sn, ppm | 1.3165 | 0.15264 | 1.0112 | 1.6217 | 0.85854 | 1.7744 | 11.59% | 23.19% | 34.78% | 1.2506 | 1.3823
Sr, ppm | 196.3146 | 8.8321 | 178.6504 | 213.9787 | 169.8183 | 222.8108 | 4.499% | 8.998% | 13.5% | 186.4988 | 206.1303
Ta, ppm | 0.24205 | 0.035028 | 0.172 | 0.31211 | 0.13697 | 0.34714 | 14.47% | 28.94% | 43.41% | 0.22995 | 0.25415
Tb, ppm | 0.50369 | 0.038635 | 0.42642 | 0.58096 | 0.38779 | 0.6196 | 7.67% | 15.34% | 23.01% | 0.47851 | 0.52888
Te, ppm | 0.15299 | 0.029414 | 0.094156 | 0.21181 | 0.064742 | 0.24123 | 19.23% | 38.45% | 57.68% | 0.14534 | 0.16063
Th, ppm | 1.4503 | 0.12857 | 1.1932 | 1.7074 | 1.0646 | 1.836 | 8.865% | 17.73% | 26.59% | 1.3778 | 1.5228
Ti, wt.% | 0.54344 | 0.01466 | 0.51412 | 0.57276 | 0.49945 | 0.58742 | 2.698% | 5.395% | 8.093% | 0.51626 | 0.57061
Tl, ppm | 0.34229 | 0.021663 | 0.29896 | 0.38561 | 0.2773 | 0.40728 | 6.329% | 12.66% | 18.99% | 0.32517 | 0.3594
Tm, ppm | 0.29292 | 0.019383 | 0.25416 | 0.33169 | 0.23477 | 0.35107 | 6.617% | 13.23% | 19.85% | 0.27827 | 0.30757
U, ppm | 0.42353 | 0.04128 | 0.34097 | 0.50609 | 0.29969 | 0.54737 | 9.747% | 19.49% | 29.24% | 0.40235 | 0.4447
V, ppm | 231.5352 | 7.0699 | 217.3954 | 245.675 | 210.3254 | 252.7449 | 3.053% | 6.107% | 9.16% | 219.9584 | 243.1119
W, ppm | 47.1217 | 3.2733 | 40.5751 | 53.6684 | 37.3018 | 56.9417 | 6.947% | 13.89% | 20.84% | 44.7656 | 49.4778
Y, ppm | 18.2682 | 1.0048 | 16.2586 | 20.2778 | 15.2538 | 21.2826 | 5.5% | 11% | 16.5% | 17.3548 | 19.1816
Yb, ppm | 1.8663 | 0.1392 | 1.5879 | 2.1447 | 1.4487 | 2.2839 | 7.459% | 14.92% | 22.38% | 1.773 | 1.9596
Zn, ppm | 173.4693 | 7.8742 | 157.7209 | 189.2177 | 149.8467 | 197.0919 | 4.539% | 9.078% | 13.62% | 164.7958 | 182.1428
Zr, ppm | 53.1248 | 4.641 | 43.8429 | 62.4067 | 39.2019 | 67.0477 | 8.736% | 17.47% | 26.21% | 50.4686 | 55.781
Aqua Regia Digestion
Ag, ppm | 1.3898 | 0.056185 | 1.2774 | 1.5021 | 1.2212 | 1.5583 | 4.043% | 8.086% | 12.13% | 1.3203 | 1.4592
Al, wt.% | 2.9899 | 0.22217 | 2.5456 | 3.4343 | 2.3234 | 3.6564 | 7.431% | 14.86% | 22.29% | 2.8404 | 3.1394
As, ppm | 95.5824 | 4.4779 | 86.6266 | 104.5382 | 82.1487 | 109.0161 | 4.685% | 9.37% | 14.05% | 90.8033 | 100.3615
B, ppm | 22.5532 | 7.2195 | 8.1142 | 36.9921 | 0.8947 | 44.2116 | 32.01% | 64.02% | 96.03% | 21.4255 | 23.6808
Ba, ppm | 50.6824 | 4.496 | 41.6904 | 59.6744 | 37.1944 | 64.1704 | 8.871% | 17.74% | 26.61% | 48.1483 | 53.2165
Be, ppm | 0.24784 | 0.041667 | 0.1645 | 0.33117 | 0.12283 | 0.37284 | 16.81% | 33.62% | 50.44% | 0.23544 | 0.26023
Bi, ppm | 0.067386 | 0.0099854 | 0.047415 | 0.087356 | 0.037429 | 0.097342 | 14.82% | 29.64% | 44.45% | 0.064016 | 0.070755
Ca, wt.% | 2.2395 | 0.2603 | 1.7189 | 2.7601 | 1.4586 | 3.0204 | 11.62% | 23.25% | 34.87% | 2.1275 | 2.3515
Cd, ppm | 0.85933 | 0.056325 | 0.74668 | 0.97198 | 0.69036 | 1.0283 | 6.555% | 13.11% | 19.66% | 0.81636 | 0.9023
Ce, ppm | 11.6936 | 0.53858 | 10.6165 | 12.7708 | 10.0779 | 13.3094 | 4.606% | 9.212% | 13.82% | 11.109 | 12.2783
Co, ppm | 29.8154 | 1.5352 | 26.7449 | 32.8858 | 25.2097 | 34.421 | 5.149% | 10.3% | 15.45% | 28.3246 | 31.3061
Cr, ppm | 34.6791 | 2.8687 | 28.9416 | 40.4165 | 26.0729 | 43.2853 | 8.272% | 16.54% | 24.82% | 32.9451 | 36.413
Cs, ppm | 0.99546 | 0.033362 | 0.92874 | 1.0622 | 0.89537 | 1.0955 | 3.351% | 6.703% | 10.05% | 0.94569 | 1.0452
Cu, ppm | 162.498 | 5.0253 | 152.4475 | 172.5486 | 147.4222 | 177.5739 | 3.093% | 6.185% | 9.278% | 154.3731 | 170.6229
Dy, ppm | 1.9331 | 0.26906 | 1.395 | 2.4712 | 1.1259 | 2.7403 | 13.92% | 27.84% | 41.76% | 1.8365 | 2.0298
Er, ppm | 1.1141 | 0.16056 | 0.79295 | 1.4352 | 0.63239 | 1.5958 | 14.41% | 28.82% | 43.24% | 1.0584 | 1.1698
Eu, ppm | 0.45028 | 0.072155 | 0.30597 | 0.59459 | 0.23381 | 0.66675 | 16.02% | 32.05% | 48.07% | 0.42777 | 0.47279
Fe, wt.% | 5.3988 | 0.15145 | 5.0959 | 5.7017 | 4.9444 | 5.8531 | 2.805% | 5.611% | 8.416% | 5.1288 | 5.6687
Ga, ppm | 9.3294 | 0.82122 | 7.687 | 10.9719 | 6.8658 | 11.7931 | 8.802% | 17.6% | 26.41% | 8.863 | 9.7959
Gd, ppm | 1.8069 | 0.19437 | 1.4182 | 2.1957 | 1.2238 | 2.3901 | 10.76% | 21.51% | 32.27% | 1.7166 | 1.8973
Ge, ppm | 0.094167 | 0.028894 | 0.036378 | 0.15196 | 0.0074834 | 0.18085 | 30.68% | 61.37% | 92.05% | 0.089458 | 0.098875
Hf, ppm | 0.43377 | 0.059154 | 0.31546 | 0.55208 | 0.25631 | 0.61123 | 13.64% | 27.27% | 40.91% | 0.41208 | 0.45546
Ho, ppm | 0.37267 | 0.0306 | 0.31147 | 0.43387 | 0.28087 | 0.46447 | 8.211% | 16.42% | 24.63% | 0.35403 | 0.3913
In, ppm | 0.038395 | 0.0034483 | 0.031498 | 0.045292 | 0.02805 | 0.04874 | 8.981% | 17.96% | 26.94% | 0.036475 | 0.040315
K, wt.% | 0.20941 | 0.012829 | 0.18375 | 0.23507 | 0.17092 | 0.2479 | 6.126% | 12.25% | 18.38% | 0.19894 | 0.21988
La, ppm | 5.4853 | 0.2705 | 4.9443 | 6.0263 | 4.6738 | 6.2968 | 4.931% | 9.863% | 14.79% | 5.211 | 5.7595
Li, ppm | 10.5696 | 0.61263 | 9.3444 | 11.7949 | 8.7318 | 12.4075 | 5.796% | 11.59% | 17.39% | 10.0412 | 11.0981
Lu, ppm | 0.12961 | 0.028541 | 0.072524 | 0.18669 | 0.043983 | 0.21523 | 22.02% | 44.04% | 66.06% | 0.12313 | 0.13609
Mg, wt.% | 1.696 | 0.084947 | 1.5261 | 1.8659 | 1.4411 | 1.9508 | 5.009% | 10.02% | 15.03% | 1.6112 | 1.7808
Mn, wt.% | 0.078831 | 0.0053418 | 0.068147 | 0.089515 | 0.062806 | 0.094857 | 6.776% | 13.55% | 20.33% | 0.07489 | 0.082773
Mo, ppm | 2.7855 | 0.1259 | 2.5337 | 3.0373 | 2.4078 | 3.1632 | 4.52% | 9.039% | 13.56% | 2.6463 | 2.9248
Na, wt.% | 0.3283 | 0.023124 | 0.28205 | 0.37455 | 0.25893 | 0.39768 | 7.044% | 14.09% | 21.13% | 0.31189 | 0.34472
Nd, ppm | 6.2576 | 0.57216 | 5.1133 | 7.4019 | 4.5411 | 7.974 | 9.143% | 18.29% | 27.43% | 5.9447 | 6.5704
Ni, ppm | 67.9061 | 3.2117 | 61.4826 | 74.3295 | 58.2709 | 77.5412 | 4.73% | 9.459% | 14.19% | 64.5108 | 71.3014
P, wt.% | 0.041627 | 0.0018336 | 0.03796 | 0.045294 | 0.036126 | 0.047128 | 4.405% | 8.81% | 13.21% | 0.039546 | 0.043708
Pb, ppm | 40.6848 | 2.0047 | 36.6754 | 44.6942 | 34.6707 | 46.699 | 4.927% | 9.855% | 14.78% | 38.6506 | 42.719
Pr, ppm | 1.392 | 0.10794 | 1.1761 | 1.6079 | 1.0682 | 1.7158 | 7.755% | 15.51% | 23.26% | 1.3224 | 1.4616
Rb, ppm | 8.9095 | 0.60759 | 7.6943 | 10.1247 | 7.0867 | 10.7323 | 6.82% | 13.64% | 20.46% | 8.464 | 9.355
Re, ppm | 0.0019375 | 0.00038472 | 0.0011681 | 0.0027069 | 0.00078333 | 0.0030917 | 19.86% | 39.71% | 59.57% | 0.0018406 | 0.0020344
S, wt.% | 0.54527 | 0.016939 | 0.51139 | 0.57914 | 0.49445 | 0.59608 | 3.107% | 6.213% | 9.32% | 0.518 | 0.57253
Sb, ppm | 1.2113 | 0.41576 | 0.37982 | 2.0429 | 0 | 2.4586 | 34.32% | 68.64% | 103% | 1.1508 | 1.2719
Sc, ppm | 5.2114 | 0.65218 | 3.907 | 6.5157 | 3.2549 | 7.1679 | 12.51% | 25.03% | 37.54% | 4.9508 | 5.472
Sm, ppm | 1.5952 | 0.17197 | 1.2512 | 1.9391 | 1.0792 | 2.1111 | 10.78% | 21.56% | 32.34% | 1.5154 | 1.6749
Sn, ppm | 0.74978 | 0.068853 | 0.61208 | 0.88749 | 0.54323 | 0.95634 | 9.183% | 18.37% | 27.55% | 0.71229 | 0.78727
Sr, ppm | 43.2605 | 6.046 | 31.1684 | 55.3526 | 25.1224 | 61.3986 | 13.98% | 27.95% | 41.93% | 41.0975 | 45.4236
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.29002 | 0.010113 | 0.2698 | 0.31025 | 0.25968 | 0.32036 | 3.487% | 6.974% | 10.46% | 0.27552 | 0.30452
Te, ppm | 0.13814 | 0.01764 | 0.10286 | 0.17342 | 0.085217 | 0.19105 | 12.77% | 25.54% | 38.31% | 0.13123 | 0.14504
Th, ppm | 1.1379 | 0.091566 | 0.95478 | 1.321 | 0.86322 | 1.4126 | 8.047% | 16.09% | 24.14% | 1.081 | 1.1948
Ti, wt.% | 0.28216 | 0.054003 | 0.17415 | 0.39017 | 0.12015 | 0.44417 | 19.14% | 38.28% | 57.42% | 0.26805 | 0.29627
Tl, ppm | 0.18027 | 0.010525 | 0.15922 | 0.20132 | 0.14869 | 0.21184 | 5.838% | 11.68% | 17.52% | 0.17126 | 0.18928
Tm, ppm | 0.14058 | 0.03233 | 0.075917 | 0.20524 | 0.043587 | 0.23757 | 23% | 46% | 68.99% | 0.13355 | 0.14761
U, ppm | 0.28384 | 0.017469 | 0.2489 | 0.31878 | 0.23143 | 0.33625 | 6.155% | 12.31% | 18.46% | 0.26965 | 0.29803
V, ppm | 115.6936 | 12.8104 | 90.0729 | 141.3144 | 77.2625 | 154.1247 | 11.07% | 22.15% | 33.22% | 109.9089 | 121.4783
W, ppm | 33.7297 | 5.4513 | 22.8271 | 44.6323 | 17.3758 | 50.0836 | 16.16% | 32.32% | 48.49% | 32.0432 | 35.4162
Y, ppm | 10.0103 | 0.55412 | 8.902 | 11.1185 | 8.3479 | 11.6726 | 5.535% | 11.07% | 16.61% | 9.5098 | 10.5108
Yb, ppm | 0.90634 | 0.11804 | 0.67025 | 1.1424 | 0.55221 | 1.2605 | 13.02% | 26.05% | 39.07% | 0.86102 | 0.95166
Zn, ppm | 158.0831 | 8.147 | 141.7891 | 174.377 | 133.6422 | 182.5239 | 5.154% | 10.31% | 15.46% | 150.1789 | 165.9872
Zr, ppm | 14.4571 | 1.9582 | 10.5407 | 18.3735 | 8.5826 | 20.3317 | 13.54% | 27.09% | 40.63% | 13.7343 | 15.18
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 240b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
N.A. | Not Applicable (Lab 1.27)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 240b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
Nagrom, Perth, WA, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 240b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 16 | 16 | 238577 | 0.08554 | 5.9046 | 5.8405
2 | 1 | ANSLu | 1 | 19 | 19 | 238578 | 0.088027 | 5.5328 | 5.8206
2 | 1 | ANSLu | 1 | 13 | 13 | 238579 | 0.087533 | 5.898 | 5.8402
2 | 1 | ANSLu | 1 | 12 | 12 | 238580 | 0.087378 | 5.6752 | 5.8282
2 | 1 | ANSLu | 1 | 14 | 14 | 238581 | 0.086978 | 5.8261 | 5.8363
2 | 1 | ANSLu | 1 | 10 | 10 | 238582 | 0.086392 | 5.7363 | 5.8315
2 | 1 | ANSLu | 1 | 1 | 1 | 238583 | 0.08626 | 5.6477 | 5.8267
2 | 1 | ANSLu | 1 | 15 | 15 | 238584 | 0.085911 | 5.723 | 5.8308
2 | 1 | ANSLu | 1 | 6 | 6 | 238585 | 0.087315 | 5.9065 | 5.8406
2 | 1 | ANSLu | 1 | 17 | 17 | 238586 | 0.087929 | 5.6861 | 5.8288
2 | 1 | ANSLu | 1 | 2 | 2 | 238587 | 0.088032 | 5.842 | 5.8372
2 | 1 | ANSLu | 1 | 7 | 7 | 238588 | 0.08745 | 5.8971 | 5.8401
2 | 1 | ANSLu | 1 | 20 | 20 | 238589 | 0.087148 | 6.1955 | 5.8561
2 | 1 | ANSLu | 1 | 9 | 9 | 238590 | 0.085972 | 5.727 | 5.831
2 | 1 | ANSLu | 1 | 5 | 5 | 238591 | 0.084098 | 5.9432 | 5.8426
2 | 1 | ANSLu | 1 | 8 | 8 | 238592 | 0.085362 | 5.9049 | 5.8405
2 | 1 | ANSLu | 1 | 11 | 11 | 238593 | 0.087628 | 5.9936 | 5.8453
2 | 1 | ANSLu | 1 | 18 | 18 | 238594 | 0.082568 | 5.8322 | 5.8366
2 | 1 | ANSLu | 1 | 4 | 4 | 238595 | 0.086942 | 5.7768 | 5.8337
2 | 1 | ANSLu | 1 | 3 | 3 | 238596 | 0.083235 | 6.0891 | 5.8504
Mean | [6 empty] | 0.086385 | 5.8369 | 5.8369
Median | [7 empty] | 5.8371 | 5.8369
SD | [7 empty] | 0.15654 | 0.0083999
RSD | [7 empty] | 2.682% | 0.1439%
RSD Upscale Factor | [7 empty] | 0.053661
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 240b (Certified Value 5.65 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 30g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 50g | 50g | 25g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 15g | 30g | 30g | 25g | 30g | 50g | 30g | 30g | 30g | 3
| 1 | 1 | 5.6477 | 4.62 | 5.82 | 5.718 | 5.5408 | 5.8706 | 5.57 | 5.66 | 5.92 | 5.42 | 5.77 | 5.611 | 6.0196 | > 5 | 5.82 | 5.7 | 5.81 | 5.857 | 5.47 | 5.69 | 5.48 | 5.627 | 5.59 | 5.879 | 5.2122 | 5.811 | 5.6424 | 5.623 | 5.39 | 5.608 | 5.578 | 5.73 | 1
| 1 | 2 | 5.842 | 4.4 | 5.795 | 5.59 | 5.6055 | 5.9396 | 5.52 | 5.81 | 5.89 | 5.52 | 5.74 | 5.653 | 6.039 | > 5 | 5.72 | 5.69 | 5.81 | 5.888 | 5.43 | 5.596 | 5.59 | 5.538 | 5.64 | 5.797 | 5.2427 | 5.602 | 5.5015 | 5.674 | 5.53 | 5.667 | 5.552 | 5.54 | #N/A
| 1 | 3 | 6.0891 | 4.419 | 5.644 | 5.608 | 5.5948 | 5.9204 | 5.54 | 5.47 | 5.97 | 5.43 | 5.76 | 5.676 | 6 | > 5 | 5.68 | 5.86 | 5.65 | 5.9 | 5.58 | 5.716 | 5.52 | 5.549 | 5.53 | 5.949 | 5.1967 | 5.559 | 5.6802 | 5.724 | 5.46 | 5.585 | 5.645 | 5.58 | 16
| 1 | 4 | 5.7768 | 5.424 | 5.742 | 5.632 | 5.5483 | 5.7416 | 5.44 | 5.53 | 5.86 | 5.35 | 5.75 | 5.659 | 6 | > 5 | 5.71 | 5.59 | 5.68 | 5.5 | 5.41 | 5.663 | 5.54 | 5.515 | 5.51 | 5.896 | 5.208 | 5.757 | 5.4511 | 5.696 | 5.36 | 5.473 | 5.712 | 5.68 | 5.6479
| 1 | 5 | 5.9432 | 5.75 | 5.703 | 5.578 | 5.5968 | 5.9004 | 5.47 | 5.69 | 5.5 | 5.4 | 5.76 | 5.689 | 6 | > 5 | 5.78 | 5.78 | 5.7 | 5.859 | 5.49 | 5.635 | 5.58 | 5.508 | 5.57 | 5.597 | 5.186 | 5.622 | 5.5028 | 5.783 | 5.51 | 5.625 | 5.66 | 5.65 | 7
| 1 | 6 | 5.9065 | 4.998 | 5.615 | 5.712 | 5.5531 | 5.7365 | 5.57 | 5.66 | 5.79 | 5.76 | 5.8 | 5.567 | 6.03 | > 5 | 5.78 | 5.71 | 5.77 | 6.061 | 5.47 | 5.641 | 5.49 | 5.502 | 5.56 | 5.623 | 5.23 | 5.779 | 5.7061 | 5.732 | 5.23 | 5.578 | 5.474 | 5.74
|  | 7 | 5.8971
|  | 8 | 5.9049
|  | 9 | 5.727
|  | 10 | 5.7363
|  | 11 | 5.9936
|  | 12 | 5.6752
|  | 13 | 5.898
|  | 14 | 5.8261
|  | 15 | 5.723
|  | 16 | 5.9046
|  | 17 | 5.6861
|  | 18 | 5.8322
|  | 19 | 5.5328
|  | 20 | 6.1955
| Mean |  | 5.8369 | 4.9352 | 5.7198 | 5.6397 | 5.5732 | 5.8515 | 5.5183 | 5.6367 | 5.8217 | 5.48 | 5.7633 | 5.6425 | 6.0148 |  | 5.7483 | 5.7217 | 5.7367 | 5.8442 | 5.475 | 5.6568 | 5.5333 | 5.5398 | 5.5667 | 5.7902 | 5.2126 | 5.6883 | 5.5807 | 5.7053 | 5.4133 | 5.5893 | 5.6035 | 5.6533
| Median |  | 5.8371 | 4.809 | 5.7225 | 5.62 | 5.574 | 5.8855 | 5.53 | 5.66 | 5.875 | 5.425 | 5.76 | 5.656 | 6.0098 |  | 5.75 | 5.705 | 5.735 | 5.8735 | 5.47 | 5.652 | 5.53 | 5.5265 | 5.565 | 5.838 | 5.2101 | 5.6895 | 5.5726 | 5.71 | 5.425 | 5.5965 | 5.6115 | 5.665
| Std Dev. |  | 0.15654 | 0.55839 | 0.081485 | 0.061154 | 0.028793 | 0.090058 | 0.053448 | 0.12094 | 0.16869 | 0.14792 | 0.020656 | 0.045518 | 0.017302 |  | 0.053072 | 0.091086 | 0.069186 | 0.18487 | 0.059245 | 0.042584 | 0.045461 | 0.046361 | 0.045898 | 0.14807 | 0.020904 | 0.10636 | 0.10822 | 0.054661 | 0.11147 | 0.06532 | 0.085743 | 0.080416
| Rel.Std.Dev. |  | 2.682% | 11.31% | 1.425% | 1.084% | 0.5166% | 1.539% | 0.9685% | 2.146% | 2.898% | 2.699% | 0.3584% | 0.8067% | 0.2877% |  | 0.9233% | 1.592% | 1.206% | 3.163% | 1.082% | 0.7528% | 0.8216% | 0.8369% | 0.8245% | 2.557% | 0.401% | 1.87% | 1.939% | 0.9581% | 2.059% | 1.169% | 1.53% | 1.422%
| PDM3 |  | 3.346% | -12.62% | 1.274% | -0.1458% | -1.322% | 3.605% | -2.294% | -0.1989% | 3.077% | -2.973% | 2.044% | -0.09559% | 6.496% |  | 1.778% | 1.306% | 1.572% | 3.475% | -3.061% | 0.1582% | -2.028% | -1.913% | -1.438% | 2.519% | -7.707% | 0.7159% | -1.19% | 1.017% | -4.153% | -1.037% | -0.7861% | 0.09623%
| Z-Score (Absolute) |  | NA | 4.84 | 0.54 | 0.01 | 0.47 | 1.44 | 0.84 | 0.03 | 1.24 | 1.1 | 0.84 | 0.01 | 2.56 | NA | 0.73 | 0.55 | 0.65 | 1.39 | 1.14 | 0.11 | 0.74 | 0.69 | 0.51 | 1.02 | 2.94 | 0.32 | 0.41 | 0.44 | 1.56 | 0.35 | 0.26 | 0.08
| Analytical results for Pd in OREAS 240b (Indicative Value 14.1 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 1
|  |  | 25g | [31 empty] | 1
| 1 | 1 | 14.8602 | [31 empty] | 1
| 1 | 2 | 14.9839 | [31 empty] | 1
| 1 | 3 | 14.3348 | [31 empty] | 16
| 1 | 4 | 14.4326 | [31 empty] | 14.1126
| 1 | 5 | 14.5892 | [31 empty] | 7
| 1 | 6 | 11.4753
| Mean |  | 14.1126
| Median |  | 14.5109
| Std Dev. |  | 1.3155
| Rel.Std.Dev. |  | 9.321%
| PDM3 |  | -2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 240b (Indicative Value 8.09 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 2
|  |  | 25g | [31 empty] | 2
| 1 | 1 | 8.0795 | [31 empty] | 1
| 1 | 2 | 8.1734 | [31 empty] | 1
| 1 | 3 | 7.9118 | [31 empty] | 16
| 1 | 4 | 7.9979 | [31 empty] | 8.0945
| 1 | 5 | 8.1752 | [31 empty] | 7
| 1 | 6 | 8.2294
| Mean |  | 8.0945
| Median |  | 8.1264
| Std Dev. |  | 0.12153
| Rel.Std.Dev. |  | 1.501%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 240b (Certified Value 5.58 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | 5.05 | 5.06 | 5.07 | 6.01 | 6.02 | 6.03 | 6.04 | ppm
|  |  | INAA | Raw*PA | Various [a] | Raw*PA | Various [b] | Various [c] | Raw*PA | Various [d] | Raw*PA | Raw*PA | Raw*PA | Various [e] | 2
|  |  | 0.085g | 400g | see footer | 400g | see footer | see footer | 400g | see footer | 450g | 400g | 450g | see footer | 3
| 1 | 1 | 5.6477 | 5.47 | 5.87 |  | 6.9393 | 5.69 | 5.3979 | 5.565 | 5.477 | 5.933 | 5.576 | 5.471 | 1
| 1 | 2 | 5.842 | 5.7 | 5.85 |  | 7.0879 | 5.606 | 5.3888 | 5.323 | 5.674 | 5.729 | 5.424 | 5.462 | #N/A
| 1 | 3 | 6.0891 | 5.74 | 5.82 |  | 6.6864 | 5.84 | 5.5865 | 5.595 | 5.594 | 5.889 | 5.609 | 5.577 | 16
| 1 | 4 | 5.7768 | 5.46 | 5.58 |  | 5.3141 | 5.44 | 5.3484 | 5.546 | 5.562 | 5.657 | 5.641 | 5.497 | 5.575
| 1 | 5 | 5.9432 | 5.56 | 5.46 |  | 5.4841 | 5.409 | 5.4548 | 5.533 | 5.695 | 5.756 | 5.578 | 5.504 | 9
| 1 | 6 | 5.9065 | 5.6 | 5.67 |  | 5.7641 | 5.531 | 5.4778 | 5.768 | 5.473 | 5.759 | 5.643 | 5.761
| 2 | 7 | 5.8971 | 5.77 | 5.64 |  | 5.55 | 5.68 | 5.526 | 5.238 | 5.814 | 5.646 | 5.6527 | 5.361
| 2 | 8 | 5.9049 | 5.63 | 5.73 |  | 5.55 | 5.5 | 5.619 | 5.523 | 5.705 | 5.69 | 5.9661 | 5.521
| 2 | 9 | 5.727 | 5.56 | 5.79 |  | 5.8 | 5.4 | 5.554 | 5.257 | 5.756 | 5.514 | 5.7645 | 5.528
| 2 | 10 | 5.7363 | 5.47 | 5.71 |  | 5.41 | 5.47 | 5.41 | 5.464 | 5.56 | 5.766 | 5.825 | 5.326
| 2 | 11 | 5.9936 | 5.76 | 5.81 |  | 5.33 | 5.48 | 5.416 | 5.374 | 5.759 | 5.896 | 5.5972 | 5.553
| 2 | 12 | 5.6752 | 5.77 | 5.91 |  | 5.3 | 5.52 | 5.425 | 5.131 | 5.719 | 5.897 | 5.8083 | 5.362
| 3 | 13 | 5.898 | 5.54 | 5.74
| 3 | 14 | 5.8261 | 5.7 | 5.52
| 3 | 15 | 5.723 | 5.47 | 5.48
| 3 | 16 | 5.9046 | 5.87 | 5.62
| 3 | 17 | 5.6861 | 5.77 | 5.64
| 3 | 18 | 5.8322 | 5.83 | 5.63
| 4 | 19 | 5.5328 | 5.66 | 5.7 | 5.39
| 4 | 20 | 6.1955 | 5.54 | 5.52 | 5.49
| 4 | 21 |  | 5.5 | 5.68 | 5.56
| 4 | 22 |  | 5.67 | 5.6 | 5.48
| 4 | 23 |  | 5.7 | 5.55
| 4 | 24 |  | 5.59 | 5.67
| 5 | 25 |  |  | 5.59 | 5.51
| 5 | 26 |  |  | 5.46 | 5.64
| 5 | 27 |  |  | 5.61 | 5.48
| 5 | 28 |  |  | 5.54 | 5.46
| 5 | 29 |  |  | 5.57
| 5 | 30 |  |  | 5.53
| 6 | 31 |  |  | 5.61 | 5.3
| 6 | 32 |  |  | 5.7 | 5.63
| 6 | 33 |  |  | 5.59 | 5.38
| 6 | 34 |  |  | 5.87 | 5.74
| 6 | 35 |  |  | 5.51
| 6 | 36 |  |  | 5.51
| Mean |  | 5.8369 | 5.6388 | 5.6467 | 5.505 | 5.8513 | 5.5472 | 5.467 | 5.4431 | 5.649 | 5.761 | 5.6737 | 5.4936
| Median |  | 5.8371 | 5.645 | 5.625 | 5.485 | 5.55 | 5.51 | 5.4399 | 5.4935 | 5.6845 | 5.7575 | 5.642 | 5.5005
| Std Dev. |  | 0.15654 | 0.12369 | 0.12412 | 0.12303 | 0.6603 | 0.13252 | 0.085943 | 0.18148 | 0.11307 | 0.12572 | 0.14407 | 0.11621
| Rel.Std.Dev. |  | 2.682% | 2.194% | 2.198% | 2.235% | 11.28% | 2.389% | 1.572% | 3.334% | 2.002% | 2.182% | 2.539% | 2.115%
| PDM3 |  | 4.697% | 1.143% | 1.285% | -1.256% | 4.956% | -0.4997% | -1.937% | -2.367% | 1.327% | 3.336% | 1.771% | -1.461%
| Z-Score (Absolute) |  | NA | 0 | 0.04 | 0.74 | 1.17 | 0.51 | 0.95 | 1.08 | 0.06 | 0.67 | 0.19 | 0.8
[a]: Lab 5.02: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (350g), Rnd 03 - Raw*PA (350g), Rnd 04 - Raw*PA (400g), Rnd 05 - Raw*PA (400g), Rnd 06 - Raw*PA (400g)
[b]: Lab 5.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[c]: Lab 5.05: Rnd 01 - Raw*PA (350g), Rnd 02 - Raw*PA (300g)
[d]: Lab 5.07: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
[e]: Lab 6.04: Rnd 01 - Raw*PA (400g), Rnd 02 - Raw*PA (300g)
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 240b (Certified Value 5.53 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.33 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.085g | 30g | 25g | 25g | 30g | 50g | 30g | 15g | 50g | 30g | 10g | 30g | 25g | 25g | 25g | 25g | 15g | 25g | 15g | 50g | 25g | 25g | 10g | 20g | 3
| 1 | 1 | 5.6477 | 5.76 | 6.12 | 4.88 | 5.5458 | 5.55 | 5.37 | 5.45 | 5.5104 | 5.15 | > 2 | 5.78 | 5.64 | 5.65 | 5.29 | 5.52 | 5.4338 | 5.451 | 5.7731 | 5.4413 | 5.49 | 4.99 | 5.475 | 5.3592 | 1
| 1 | 2 | 5.842 | 5.65 | 6.08 | 4.88 | 5.565 | 5.5 | 5.09 | 5.35 | 5.3312 | 5.36 | > 2 | 5.78 | 5.46 | 5.71 | 5.36 | 5.47 | 5.5208 | 5.563 | 5.5553 | 5.3358 | 5.53 | 4.85 | 5.519 | 5.5313 | #N/A
| 1 | 3 | 6.0891 | 5.82 | 6.06 | 4.8 | 5.6076 | 5.49 | 5.33 | 5.93 | 5.3984 | 5.26 | > 2 | 5.78 | 5.4 | 5.54 | 5.35 | 5.46 | 5.6135 | 5.478 | 5.5148 | 5.5486 | 5.42 | 5.06 | 5.586 | 5.3913 | 16
| 1 | 4 | 5.7768 | 5.92 | 6.2 | 4.8 | 5.6886 | 5.35 | 5.07 | 5.77 | 5.5328 | 5.33 | > 2 | 5.8 | 5.37 | 5.66 | 5.5 | 5.35 | 5.5826 | 5.528 | 4.8469 | 5.4604 | 5.66 | 4.98 | 5.561 | 5.599 | 5.5261
| 1 | 5 | 5.9432 | 5.89 | 6.39 | 4.72 | 5.5456 | 5.43 | 5.25 | 5.86 | 5.4216 | 5.36 | > 2 | 5.86 | 5.37 | 5.71 | 5.28 | 5.35 | 5.7062 | 5.608 | 5.6041 | 5.4797 | 5.49 | 4.88 | 5.624 | 5.4144 | 11
| 1 | 6 | 5.9065 | 5.83 | 6.12 | 4.56 | 5.5266 | 5.49 | 5.34 | 6.13 | 5.5296 | 5.39 | > 2 | 5.8152 | 5.63 | 5.64 | 5.2 | 5.44 | 5.4796 | 5.564 | 5.7245 | 5.3634 | 5.595 | 5.04 | 5.607 | 5.5868
|  | 7 | 5.8971
|  | 8 | 5.9049
|  | 9 | 5.727
|  | 10 | 5.7363
|  | 11 | 5.9936
|  | 12 | 5.6752
|  | 13 | 5.898
|  | 14 | 5.8261
|  | 15 | 5.723
|  | 16 | 5.9046
|  | 17 | 5.6861
|  | 18 | 5.8322
|  | 19 | 5.5328
|  | 20 | 6.1955
| Mean |  | 5.8369 | 5.8117 | 6.1617 | 4.7733 | 5.5799 | 5.4683 | 5.2417 | 5.7483 | 5.454 | 5.3083 |  | 5.8025 | 5.4783 | 5.6517 | 5.33 | 5.4317 | 5.5561 | 5.532 | 5.5031 | 5.4382 | 5.5308 | 4.9667 | 5.562 | 5.4803
| Median |  | 5.8371 | 5.825 | 6.12 | 4.8 | 5.5554 | 5.49 | 5.29 | 5.815 | 5.466 | 5.345 |  | 5.79 | 5.43 | 5.655 | 5.32 | 5.45 | 5.5517 | 5.5455 | 5.5797 | 5.4508 | 5.51 | 4.985 | 5.5735 | 5.4728
| Std Dev. |  | 0.15654 | 0.09704 | 0.12172 | 0.12044 | 0.059975 | 0.069402 | 0.13152 | 0.29641 | 0.082857 | 0.089312 |  | 0.031599 | 0.12576 | 0.062423 | 0.10119 | 0.068532 | 0.098572 | 0.058744 | 0.33633 | 0.078094 | 0.085347 | 0.084774 | 0.056328 | 0.10484
| Rel.Std.Dev. |  | 2.682% | 1.67% | 1.976% | 2.523% | 1.075% | 1.269% | 2.509% | 5.156% | 1.519% | 1.682% |  | 0.5446% | 2.296% | 1.105% | 1.899% | 1.262% | 1.774% | 1.062% | 6.112% | 1.436% | 1.543% | 1.707% | 1.013% | 1.913%
| PDM3 |  | 5.624% | 5.168% | 11.5% | -13.62% | 0.9733% | -1.045% | -5.147% | 4.022% | -1.304% | -3.94% |  | 5.003% | -0.864% | 2.273% | -3.548% | -1.709% | 0.5429% | 0.1071% | -0.4156% | -1.59% | 0.08599% | -10.12% | 0.65% | -0.8279%
| Z-Score (Absolute) |  | NA | 2.72 | 5.7 | 6.12 | 0.75 | 0.2 | 2.13 | 2.18 | 0.32 | 1.56 | NA | 2.65 | 0.11 | 1.36 | 1.38 | 0.51 | 0.55 | 0.34 | 0.1 | 0.46 | 0.33 | 4.47 | 0.6 | 0.1
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 240b (Certified Value 5.44 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.30 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 05g | 50g | 60g | 30g | 30g | 30g | 200g | 30g | 200g | 200g | 20g | 30g | 05g | 3
| 1 | 1 | 5.6477 | 4.484 | 5.62 | 4.2 | 5.308 | 5.0064 | 4.98 | 5.77 | 5.28 | 5.4 | 5.3076 | 5.45 | 5.88 | 5.56 | 5.54 | 5.43 | 5.51 | 5.08 | 5.3866 | 5.35 | 5.528 | 5.6 | 1
| 1 | 2 | 5.842 | 4.562 | 5.57 | 4.225 | 5.316 | 4.9472 | 4.58 | 5.73 | 5.26 | 5.4 | 5.204 | 5.62 | 5.8 | 5.46 | 5.54 | 5.44 | 5.45 | 5.1 | 5.3957 | 5.26 | 5.435 | 5.58 | #N/A
| 1 | 3 | 6.0891 | 4.267 | 5.59 | 4.475 | 5.392 | 4.9437 | 5.12 | 5.65 | 5.44 | 5.3 | 5.2956 | 5.54 | 5.83 | 5.44 | 5.64 | 5.414 | 5.5 | 5.07 | 5.4215 | 5.35 | 5.447 | 5.56 | 16
| 1 | 4 | 5.7768 | 4.203 | 5.55 | 4.075 | 5.696 | 5.0531 | 4.92 | 5.95 | 5.39 | 5.3 | 5.3616 | 5.55 | 5.79 | 5.51 | 5.46 | 5.39 | 5.54 | 5.19 | 5.543 | 5.3 | 5.482 | 5.47 | 5.4353
| 1 | 5 | 5.9432 | 4.549 | 5.4 | 4.225 | 5.424 | 5.0008 | 4.93 | 5.69 | 5.41 | 5.5 | 5.254 | 5.54 | 5.83 | 5.33 | 5.31 | 5.393 | 5.46 | 5.16 | 5.5002 | 5.42 | 5.433 | 5.62 | 13
| 1 | 6 | 5.9065 | 4.51 | 5.41 | 5.175 | 5.388 | 5.0421 | 4.88 | 5.62 | 5.56 | 5.2 | 5.288 | 5.51 | 5.81 | 5.38 | 5.43 | 5.424 | 5.46 | 5.09 | 5.5376 | 5.34 | 5.425 | 5.53
|  | 7 | 5.8971
|  | 8 | 5.9049
|  | 9 | 5.727
|  | 10 | 5.7363
|  | 11 | 5.9936
|  | 12 | 5.6752
|  | 13 | 5.898
|  | 14 | 5.8261
|  | 15 | 5.723
|  | 16 | 5.9046
|  | 17 | 5.6861
|  | 18 | 5.8322
|  | 19 | 5.5328
|  | 20 | 6.1955
| Mean |  | 5.8369 | 4.4292 | 5.5233 | 4.3958 | 5.4207 | 4.9989 | 4.9017 | 5.735 | 5.39 | 5.35 | 5.2852 | 5.535 | 5.8233 | 5.4467 | 5.4867 | 5.4152 | 5.4867 | 5.115 | 5.4641 | 5.3367 | 5.4583 | 5.56
| Median |  | 5.8371 | 4.497 | 5.56 | 4.225 | 5.39 | 5.0036 | 4.925 | 5.71 | 5.4 | 5.35 | 5.2918 | 5.54 | 5.82 | 5.45 | 5.5 | 5.419 | 5.48 | 5.095 | 5.4608 | 5.345 | 5.441 | 5.57
| Std Dev. |  | 0.15654 | 0.15427 | 0.094587 | 0.40324 | 0.14238 | 0.046011 | 0.17826 | 0.11828 | 0.11027 | 0.10488 | 0.052974 | 0.055408 | 0.032042 | 0.083825 | 0.11343 | 0.020203 | 0.03559 | 0.048477 | 0.071323 | 0.053914 | 0.039596 | 0.054037
| Rel.Std.Dev. |  | 2.682% | 3.483% | 1.712% | 9.173% | 2.627% | 0.9204% | 3.637% | 2.062% | 2.046% | 1.96% | 1.002% | 1.001% | 0.5502% | 1.539% | 2.067% | 0.3731% | 0.6487% | 0.9477% | 1.305% | 1.01% | 0.7254% | 0.9719%
| PDM3 |  | 7.388% | -18.51% | 1.619% | -19.12% | -0.2699% | -8.03% | -9.819% | 5.513% | -0.8341% | -1.57% | -2.763% | 1.834% | 7.138% | 0.2084% | 0.9444% | -0.3711% | 0.9444% | -5.894% | 0.5288% | -1.815% | 0.4231% | 2.294%
| Z-Score (Absolute) |  | NA | 6.51 | 0.67 | 6.73 | 0 | 2.77 | 3.41 | 2.06 | 0.2 | 0.46 | 0.89 | 0.75 | 2.64 | 0.17 | 0.43 | 0.04 | 0.43 | 2.01 | 0.29 | 0.55 | 0.25 | 0.92
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 240b (Certified Value 1.4 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.95 | 1.4 | 1.3793 | 1.4 | 1.34 | 1.4 | 1.43 | 1.5 | 1.3 | 1.39 | 1.46 | 1.37 | 1.45 | 1.5 | 1.2 | 1.3 | 1.38 | 0.29 | 2 | 1.39 | 1.3423 | 1.43 | 1.45 | 1.41 | 1.42 |  |  |  |  | 1
| 1 | 2 | 2.05 | 1.4 | 1.3807 | 1.3 | 1.56 | 1.4 | 1.46 | 1.5 | 1.4 | 1.45 | 1.48 | 1.48 | 1.39 | 1.4 | 1.3 | 1.3 | 1.42 | 0.28 | 2 | 1.37 | 1.384 | 1.34 | 1.42 | 1.41 | 1.36 |  |  |  |  | 18
| 1 | 3 | 2.02 | 1.35 | 1.4049 | 1.4 | 1.56 | 1.4 | 1.46 | 1.5 | 1.3 | 1.39 | 1.53 | 1.4 | 1.4 | 1.5 | 1.3 | 1.3 | 1.36 | 0.24 | 2 | 1.48 | 1.3565 | 1.45 | 1.46 | 1.39 | 1.27 |  |  |  |  | 16
| 1 | 4 | 2.01 | 1.39 | 1.426 | 1.4 | 1.53 | 1.4 | 1.46 | 1.5 | 1.4 | 1.39 | 1.45 | 1.41 | 1.44 | 1.4 | 1.2 | 1.2 | 1.38 | 0.34 | 2 | 1.42 | 1.3703 | 1.37 | 1.38 | 1.42 | 1.39 |  |  |  |  | 1.3996
| 1 | 5 | 1.97 | 1.36 | 1.3916 | 1.5 | 1.5 | 1.4 | 1.38 | 1.5 | 1.4 | 1.4 | 1.44 | 1.39 | 1.39 | 1.3 | 1.2 | 1.2 | 1.39 | 0.28 | 2 | 1.44 | 1.3033 | 1.41 | 1.4 | 1.35 | 1.26 |  |  |  |  | 15
| 1 | 6 | 2.03 | 1.43 | 1.4071 | 1.5 | 1.7 | 1.4 | 1.38 | 1.5 | 1.4 | 1.46 | 1.39 | 1.41 | 1.37 | 1.4 | 1.3 | 1.3 | 1.37 | 0.26 | 2 | 1.44 | 1.2711 | 1.39 | 1.43 | 1.37 | 1.37
| Mean |  | 2.005 | 1.3883 | 1.3983 | 1.4167 | 1.5317 | 1.4 | 1.4283 | 1.5 | 1.3667 | 1.4133 | 1.4583 | 1.41 | 1.4067 | 1.4167 | 1.25 | 1.2667 | 1.3833 | 0.28167 | 2 | 1.4233 | 1.3379 | 1.3983 | 1.4233 | 1.3917 | 1.345
| Median |  | 2.015 | 1.395 | 1.3983 | 1.4 | 1.545 | 1.4 | 1.445 | 1.5 | 1.4 | 1.395 | 1.455 | 1.405 | 1.395 | 1.4 | 1.25 | 1.3 | 1.38 | 0.28 | 2 | 1.43 | 1.3494 | 1.4 | 1.425 | 1.4 | 1.365
| Std Dev. |  | 0.037815 | 0.029269 | 0.017917 | 0.075277 | 0.11635 | 2.4324e-16 | 0.0392 | 0 | 0.05164 | 0.03266 | 0.046224 | 0.037417 | 0.031411 | 0.075277 | 0.054772 | 0.05164 | 0.020656 | 0.033714 | 0 | 0.039328 | 0.042894 | 0.040208 | 0.030111 | 0.027142 | 0.065345
| Rel.Std.Dev. |  | 1.886% | 2.108% | 1.281% | 5.314% | 7.596% | 1.737e-14% | 2.744% | 0% | 3.779% | 2.311% | 3.17% | 2.654% | 2.233% | 5.314% | 4.382% | 4.077% | 1.493% | 11.97% | 0% | 2.763% | 3.206% | 2.875% | 2.116% | 1.95% | 4.858%
| PDM3 |  | 43.26% | -0.8026% | -0.09248% | 1.222% | 9.439% | 0.03101% | 2.055% | 7.176% | -2.351% | 0.9837% | 4.199% | 0.7455% | 0.5073% | 1.222% | -10.69% | -9.496% | -1.16% | -79.87% | 42.9% | 1.698% | -4.405% | -0.08808% | 1.698% | -0.5644% | -3.899%
| Z-Score (Absolute) |  | 20.29 | 0.51 | 0.17 | 0.45 | 4.33 | 0.11 | 0.84 | 3.26 | 1.24 | 0.34 | 1.85 | 0.22 | 0.11 | 0.45 | 5.17 | 4.61 | 0.67 | 37.82 | NA | 0.67 | 2.21 | 0.17 | 0.67 | 0.39 | 1.97
| Analytical results for Al in OREAS 240b (Certified Value 6.49 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 7.01 | 6.2387 | 6.7518 | 6.5021 | 6.35 | 7.092 | 6.7 | 6.082 | 6.2 | 6.46 | 6.5 | 6.44 | 6.32 | 6.23 | 6.69 | 6.476 | 6.51 | 6.4 | 4.15 | 6.2579 | 6.409 | 6.39 | 6.048 | 6.47 | 7.02 | 6.68 | 7.5 |  |  | 1
| 1 | 2 | 7.09 | 6.2295 | 6.8236 | 6.518 | 6.33 | 7.092 | 6.71 | 6.158 | 6.25 | 6.53 | 6.49 | 6.75 | 6.49 | 6.14 | 6.45 | 6.608 | 6.43 | 6.33 | 3.4 | 6.2851 | 6.419 | 6.56 | 6.149 | 6.63 | 6.78 | 6.68 | 7 |  |  | #N/A
| 1 | 3 | 7.08 | 6.3373 | 6.8522 | 6.5674 | 6.31 | 7.039 | 6.73 | 6.4399 | 6.17 | 6.42 | 6.41 | 6.72 | 6.27 | 6.15 | 6.64 | 6.603 | 6.44 | 6.18 | 4.51 | 6.4005 | 6.347 | 6.51 | 6.195 | 6.5 | 6.99 | 6.72 | 6 |  |  | 16
| 1 | 4 | 6.94 | 6.2174 | 6.7763 | 6.4554 | 6.37 | 7.198 | 6.76 | 6.2539 | 6.14 | 6.25 | 6.4 | 6.58 | 6.34 | 6.28 | 6.6 | 6.592 | 6.36 | 6.33 | 3.69 | 6.4086 | 6.369 | 6.51 | 6.242 | 6.62 | 6.98 | 6.69 | 7.4 |  |  | 6.4931
| 1 | 5 | 6.33 | 6.0691 | 6.7878 | 6.3639 | 6.36 | 7.039 | 6.62 | 6.0156 | 6.19 | 6.61 | 6.61 | 6.66 | 6.1 | 6.26 | 6.54 | 6.475 | 6.36 | 6.54 | 4.58 | 6.5156 | 6.392 | 6.48 | 5.974 | 6.63 | 6.83 | 6.65 | 5.7 |  |  | 16
| 1 | 6 | 6.7 | 6.3873 | 6.7848 | 6.4204 | 6.32 | 7.145 | 6.59 | 6.1025 | 6.21 | 6.64 | 6.59 | 6.35 | 6.16 | 6.22 | 6.32 | 6.598 | 6.29 | 6.34 | 4.03 | 6.4637 | 6.422 | 6.54 | 6.184 | 6.52 | 6.72 | 6.63 | 6.9
| Mean |  | 6.8583 | 6.2465 | 6.7961 | 6.4712 | 6.34 | 7.1008 | 6.685 | 6.1753 | 6.1933 | 6.485 | 6.5 | 6.5833 | 6.28 | 6.2133 | 6.54 | 6.5587 | 6.3983 | 6.3533 | 4.06 | 6.3886 | 6.393 | 6.4983 | 6.132 | 6.5617 | 6.8867 | 6.675 | 6.75
| Median |  | 6.975 | 6.2341 | 6.7863 | 6.4788 | 6.34 | 7.092 | 6.705 | 6.1303 | 6.195 | 6.495 | 6.495 | 6.62 | 6.295 | 6.225 | 6.57 | 6.595 | 6.395 | 6.335 | 4.09 | 6.4046 | 6.4005 | 6.51 | 6.1665 | 6.57 | 6.905 | 6.68 | 6.95
| Std Dev. |  | 0.29553 | 0.11021 | 0.035948 | 0.073134 | 0.023664 | 0.061959 | 0.065955 | 0.15231 | 0.037238 | 0.14265 | 0.087636 | 0.15958 | 0.13871 | 0.057155 | 0.13609 | 0.06464 | 0.077309 | 0.1169 | 0.45922 | 0.10011 | 0.029886 | 0.059805 | 0.10108 | 0.073052 | 0.12612 | 0.031464 | 0.73959
| Rel.Std.Dev. |  | 4.309% | 1.764% | 0.529% | 1.13% | 0.3733% | 0.8726% | 0.9866% | 2.467% | 0.6013% | 2.2% | 1.348% | 2.424% | 2.209% | 0.9199% | 2.081% | 0.9856% | 1.208% | 1.84% | 11.31% | 1.567% | 0.4675% | 0.9203% | 1.648% | 1.113% | 1.831% | 0.4714% | 10.96%
| PDM3 |  | 5.626% | -3.796% | 4.667% | -0.3362% | -2.357% | 9.36% | 2.956% | -4.893% | -4.616% | -0.124% | 0.107% | 1.39% | -3.281% | -4.308% | 0.723% | 1.011% | -1.459% | -2.152% | -37.47% | -1.609% | -1.541% | 0.08133% | -5.561% | 1.057% | 6.062% | 2.802% | 3.957%
| Z-Score (Absolute) |  | 1.32 | 0.85 | 1.1 | 0.05 | 0.51 | 2.18 | 0.71 | 1.1 | 1.04 | 0 | 0.05 | 0.35 | 0.73 | 0.96 | 0.2 | 0.26 | 0.31 | 0.47 | 8.61 | 0.34 | 0.33 | 0.05 | 1.25 | 0.27 | 1.43 | 0.67 | 0.94
| Analytical results for As in OREAS 240b (Certified Value 98 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 102 | 97.1 | 102.4804 | 78.3 | 100 | 95.4 | 92 | 83.5 | 99.5 | 99.9 | 97.6 | 97.1 | 103 | 97 | 102 | 97 | 29.53 | 95 | 98.3 | 87.6049 | 102 | 103 | 89.9 | 101 | [5 empty] | 1
| 1 | 2 | 103 | 99.7 | 103.0218 | 71.1 | 100 | 94 | 90 | 85.1 | 101.5 | 103 | 100.5 | 96.8 | 99 | 102 | 100 | 98 | 47.507 | 92 | 96.6 | 93.9313 | 100 | 100 | 87 | 84.7 | [5 empty] | 19
| 1 | 3 | 108 | 98.8 | 103.6116 | 76.6 | 102 | 98.2 | 89 | 82 | 98.7 | 100 | 98 | 99.5 | 102 | 96 | 102 | 96 | 38.46 | 94 | 101 | 90.0992 | 102 | 104 | 89.1 | 87.4 | [5 empty] | 16
| 1 | 4 | 101 | 98.9 | 102.3668 | 89.9 | 102 | 95.3 | 90 | 81.3 | 100.5 | 99.2 | 99.2 | 100 | 102 | 96 | 99 | 94 | 55.6 | 92 | 100 | 87.3087 | 101 | 99 | 91.4 | 100 | [5 empty] | 97.988
| 1 | 5 | 107 | 96.1 | 101.5568 | 101 | 98 | 93.1 | 91 | 89.3 | 101 | 98.6 | 96.7 | 97.6 | 100 | 99 | 97 | 98 | 54.65 | 95 | 100 | 84.7848 | 105 | 98 | 85.2 | 89.4 | [5 empty] | 17
| 1 | 6 | 102 | 98.8 | 101.5184 | 93.4 | 99 | 93.8 | 92 | 92.9 | 102 | 102 | 97.9 | 97.4 | 97 | 101 | 98 | 97 | 50.52 | 93 | 99.5 | 78.963 | 104 | 99 | 85.5 | 96.6
| Mean |  | 103.8333 | 98.2333 | 102.426 | 85.05 | 100.1667 | 94.9667 | 90.6667 | 85.6833 | 100.5333 | 100.45 | 98.3167 | 98.0667 | 100.5 | 98.5 | 99.6667 | 96.6667 | 46.0445 | 93.5 | 99.2333 | 87.1153 | 102.3333 | 100.5 | 88.0167 | 93.1833
| Median |  | 102.5 | 98.8 | 102.4236 | 84.1 | 100 | 94.65 | 90.5 | 84.3 | 100.75 | 99.95 | 97.95 | 97.5 | 101 | 98 | 99.5 | 97 | 49.0135 | 93.5 | 99.75 | 87.4568 | 102 | 99.5 | 88.05 | 93
| Std Dev. |  | 2.9269 | 1.3471 | 0.81806 | 11.4823 | 1.6021 | 1.8184 | 1.2111 | 4.5398 | 1.2437 | 1.6968 | 1.3378 | 1.3411 | 2.2583 | 2.5884 | 2.0656 | 1.5055 | 10.1763 | 1.3784 | 1.5603 | 5.0503 | 1.8619 | 2.429 | 2.5087 | 6.9133
| Rel.Std.Dev. |  | 2.819% | 1.371% | 0.7987% | 13.5% | 1.599% | 1.915% | 1.336% | 5.298% | 1.237% | 1.689% | 1.361% | 1.368% | 2.247% | 2.628% | 2.072% | 1.557% | 22.1% | 1.474% | 1.572% | 5.797% | 1.819% | 2.417% | 2.85% | 7.419%
| PDM3 |  | 5.965% | 0.2504% | 4.529% | -13.2% | 2.223% | -3.083% | -7.472% | -12.56% | 2.598% | 2.513% | 0.3355% | 0.08032% | 2.564% | 0.5225% | 1.713% | -1.348% | -53.01% | -4.58% | 1.271% | -11.1% | 4.435% | 2.564% | -10.18% | -4.903%
| Z-Score (Absolute) |  | 1.35 | 0.01 | 1.01 | 3.2 | 0.46 | 0.8 | 1.84 | 3.05 | 0.55 | 0.53 | 0.01 | 0.05 | 0.54 | 0.05 | 0.34 | 0.39 | 12.65 | 1.16 | 0.23 | 2.7 | 0.98 | 0.54 | 2.49 | 1.23
| Analytical results for B in OREAS 240b (Indicative Value 33.8 ppm) | [30 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.03 | 1.25 | 1.27 | [26 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 1
1
| 1 | 1 | 58.48 | 14684.37 | 10 | [26 empty] | 1
| 1 | 2 | 54.48 | 14900.82 | 10 | [26 empty] | 3
| 1 | 3 | 59.06 | 10290.02 | 10 | [26 empty] | 16
| 1 | 4 | 55.28 | 19537.41 | 10 | [26 empty] | 33.7917
| 1 | 5 | 58.5 | 16945.14 | 10 | [26 empty] | 9
| 1 | 6 | 59.7 | 14833.53 | 10
| Mean |  | 57.5833 | 15198.5483 | 10
| Median |  | 58.49 | 14867.175 | 10
| Std Dev. |  | 2.156 | 3048.154 | 0
| Rel.Std.Dev. |  | 3.744% | 20.06% | 0%
| PDM3 |  | 70.41% | 44877% | -70.41%
| Z-Score (Absolute) |  | 0 | 214.56 | 0.67
| Analytical results for Ba in OREAS 240b (Certified Value 324 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 324 | 321.1 | 340.9506 | 341.5372 | 298 | 362 | 321 | 319.4 | 307 | 324 | 330 | 320 | 350 | 320 | 270 | 315 | 295 | 298 | 1886.51 | 304.587 | 320 | 319.5716 | 320 | 343.7671 | 324 | 356 | 315 | 369 |  | 1
| 1 | 2 | 329 | 326 | 341.8158 | 324.0198 | 297 | 365 | 327 | 318.3 | 306 | 322 | 330 | 340 | 340 | 320 | 263 | 322 | 305 | 295 | 1880.42 | 300.267 | 320 | 314.9613 | 330 | 346.5411 | 331 | 361 | 312 | 352 |  | 20
| 1 | 3 | 331 | 329.6 | 342.4956 | 328.118 | 301 | 362 | 327 | 327.9 | 305 | 336 | 320 | 330 | 320 | 320 | 265 | 321 | 298 | 288 | 1281.03 | 310.707 | 320 | 321.3809 | 330 | 346.2318 | 325 | 379 | 316 | 333 |  | 16
| 1 | 4 | 325 | 327.8 | 338.9112 | 326.0802 | 298 | 367 | 327 | 320.1 | 307 | 326 | 320 | 330 | 320 | 320 | 264 | 320 | 289 | 312 | 1836.75 | 309.069 | 320 | 324.6438 | 330 | 343.7462 | 331 | 356 | 312 | 380 |  | 323.8994
| 1 | 5 | 316 | 326.5 | 334.4719 | 324.6938 | 320 | 368 | 327 | 309.3 | 308 | 331 | 330 | 340 | 310 | 320 | 259 | 315 | 294 | 305 | 1912.74 | 315.378 | 315 | 325.6032 | 330 | 330.5897 | 332 | 341 | 312 | 323 |  | 18
| 1 | 6 | 326 | 323.7 | 341.1463 | 322.2168 | 296 | 366 | 321 | 315.4 | 307 | 333 | 330 | 320 | 310 | 320 | 253 | 321 | 295 | 296 | 1469.07 | 313.317 | 320 | 333.3475 | 330 | 316.6885 | 326 | 357 | 309 | 360
| Mean |  | 325.1667 | 325.7833 | 339.9652 | 327.7776 | 301.6667 | 365 | 325 | 318.4 | 306.6667 | 328.6667 | 326.6667 | 330 | 325 | 320 | 262.3333 | 319 | 296 | 299 | 1711.0867 | 308.8875 | 319.1667 | 323.2514 | 328.3333 | 337.9274 | 328.1667 | 358.3333 | 312.6667 | 352.8333
| Median |  | 325.5 | 326.25 | 341.0485 | 325.387 | 298 | 365.5 | 327 | 318.85 | 307 | 328.5 | 330 | 330 | 320 | 320 | 263.5 | 320.5 | 295 | 297 | 1858.585 | 309.888 | 320 | 323.0123 | 330 | 343.7567 | 328.5 | 356.5 | 312 | 356
| Std Dev. |  | 5.1929 | 3.0156 | 2.9489 | 7.0268 | 9.136 | 2.5298 | 3.0984 | 6.0985 | 1.0328 | 5.5015 | 5.164 | 8.9443 | 16.4317 | 0 | 5.785 | 3.1623 | 5.2915 | 8.3905 | 268.1131 | 5.6196 | 2.0412 | 6.2477 | 4.0825 | 11.9671 | 3.5449 | 12.2257 | 2.5033 | 21.6094
| Rel.Std.Dev. |  | 1.597% | 0.9256% | 0.8674% | 2.144% | 3.029% | 0.6931% | 0.9533% | 1.915% | 0.3368% | 1.674% | 1.581% | 2.71% | 5.056% | 0% | 2.205% | 0.9913% | 1.788% | 2.806% | 15.67% | 1.819% | 0.6396% | 1.933% | 1.243% | 3.541% | 1.08% | 3.412% | 0.8006% | 6.125%
| PDM3 |  | 0.3912% | 0.5816% | 4.96% | 1.197% | -6.864% | 12.69% | 0.3398% | -1.698% | -5.32% | 1.472% | 0.8544% | 1.883% | 0.3398% | -1.204% | -19.01% | -1.513% | -8.614% | -7.687% | 428.3% | -4.635% | -1.461% | -0.2001% | 1.369% | 4.331% | 1.317% | 10.63% | -3.468% | 8.933%
| Z-Score (Absolute) |  | 0.01 | 0.07 | 1.57 | 0.28 | 2.47 | 4.22 | 0.01 | 0.71 | 1.95 | 0.38 | 0.17 | 0.52 | 0.01 | 0.54 | 6.63 | 0.64 | 3.07 | 2.76 | 146.41 | 1.71 | 0.63 | 0.19 | 0.34 | 1.36 | 0.33 | 3.51 | 1.31 | 2.93
| Analytical results for Be in OREAS 240b (Certified Value 0.47 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.5 | 0.52 | 0.44 | 0.45964 | < 0.5 | 0.5 | 0.5 | 0.7 | 0.5 | < 5 | < 1 | 0.49 | 0.49 | 0.53 | 0.46 | < 0.5 | 0.5 | < 1 | 0.46 | < 0.5 | 0.41907 | 0.48 | 0.39139 | 0.18 | 0.6 | 0.47 | 0.55 |  |  | 1
| 1 | 2 | 0.5 | 0.55 | 0.41 | 0.46024 | < 0.5 | 0.5 | 0.6 | 0.6 | 0.45 | < 5 | < 1 | 0.5 | 0.53 | 0.52 | 0.46 | < 0.5 | 0.5 | < 1 | 0.5 | 0.5 | 0.38589 | 0.46 | 0.39586 | 0.17 | 0.7 | 0.45 | 0.48 |  |  | 21
| 1 | 3 | 0.6 | 0.54 | 0.4 | 0.4527 | < 0.5 | 0.5 | 0.6 | 0.6 | 0.46 | < 5 | < 1 | 0.48 | 0.53 | 0.52 | 0.46 | < 0.5 | 0.5 | < 1 | 0.47 | < 0.5 | 0.47062 | 0.48 | 0.39674 | 0.14 | 0.6 | 0.47 | 0.48 |  |  | 16
| 1 | 4 | 0.5 | 0.51 | 0.37 | 0.45327 | 0.50214 | 0.5 | 0.6 | 0.5 | 0.52 | < 5 | < 1 | 0.49 | 0.52 | 0.52 | 0.47 | < 0.5 | 0.5 | < 1 | 0.49 | < 0.5 | 0.40944 | 0.48 | 0.39975 | 0.16 | 0.6 | 0.47 | 0.68 |  |  | 0.47001
| 1 | 5 | 0.4 | 0.52 | 0.37 | 0.45368 | < 0.5 | 0.5 | 0.6 | 0.5 | 0.45 | < 5 | < 1 | 0.51 | 0.53 | 0.51 | 0.46 | < 0.5 | 0.5 | < 1 | 0.45 | < 0.5 | 0.43746 | 0.48 | 0.38944 | 0.17 | 0.6 | 0.42 | 0.33 |  |  | 19
| 1 | 6 | 0.5 | 0.55 | 0.36 | 0.46969 | < 0.5 | 0.5 | 0.6 | 0.7 | 0.47 | < 5 | < 1 | 0.49 | 0.46 | 0.52 | 0.47 | < 0.5 | 0.5 | < 1 | 0.49 | 0.5 | 0.46267 | 0.49 | 0.38651 | 0.15 | 0.6 | 0.43 | 0.48
| Mean |  | 0.5 | 0.53167 | 0.39167 | 0.4582 | 0.50214 | 0.5 | 0.58333 | 0.6 | 0.475 |  |  | 0.49333 | 0.51 | 0.52 | 0.46333 |  | 0.5 |  | 0.47667 | 0.5 | 0.43086 | 0.47833 | 0.39328 | 0.16167 | 0.61667 | 0.45167 | 0.5
| Median |  | 0.5 | 0.53 | 0.385 | 0.45666 | 0.50214 | 0.5 | 0.6 | 0.6 | 0.465 |  |  | 0.49 | 0.525 | 0.52 | 0.46 |  | 0.5 |  | 0.48 | 0.5 | 0.42827 | 0.48 | 0.39362 | 0.165 | 0.6 | 0.46 | 0.48
| Std Dev. |  | 0.063246 | 0.017224 | 0.030605 | 0.0065291 |  | 0 | 0.040825 | 0.089443 | 0.02881 |  |  | 0.010328 | 0.028983 | 0.0063246 | 0.005164 |  | 0 |  | 0.019664 | 0 | 0.032423 | 0.0098319 | 0.0049935 | 0.01472 | 0.040825 | 0.022286 | 0.11402
| Rel.Std.Dev. |  | 12.65% | 3.24% | 7.814% | 1.425% |  | 0% | 6.999% | 14.91% | 6.065% |  |  | 2.094% | 5.683% | 1.216% | 1.115% |  | 0% |  | 4.125% | 0% | 7.525% | 2.055% | 1.27% | 9.105% | 6.62% | 4.934% | 22.8%
| PDM3 |  | 6.381% | 13.12% | -16.67% | -2.512% | 6.835% | 6.381% | 24.11% | 27.66% | 1.062% |  |  | 4.962% | 8.508% | 10.64% | -1.421% |  | 6.381% |  | 1.416% | 6.381% | -8.33% | 1.771% | -16.32% | -65.6% | 31.2% | -3.903% | 6.381%
| Z-Score (Absolute) |  | NA | 1.11 | 1.72 | 0.37 | 3.74 | NA | NA | 2.49 | 0.03 | 40.93 | 0.47 | 0.34 | 0.67 | 0.88 | 0.27 | 4.59 | NA | 0.47 | 0 | NA | 0.93 | 0.03 | 1.69 | 6.37 | NA | 0.51 | 0.47
Results from laboratories 1.01, 1.07, 1.08, 1.22, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 240b (Certified Value 0.073 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.1 | 0.08 | 0.092736 | 0.07 | 0.2 | 0.09 | 0.07 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.06 | < 5 | 0.06 | < 0.1 | 0.08 | < 0.1 | < 0.1 | < 0.11 | 0.07 | 0.0502 | 0.07 | 0.09 | 0.07 | 0.06 |  |  |  |  | 1
| 1 | 2 | 0.1 | 0.07 | 0.091014 | 0.07 | 0.3 | 0.09 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.08 | 0.07 | < 5 | 0.08 | < 0.1 | 0.07 | < 0.1 | < 0.1 | < 0.11 | 0.07 | 0.0467 | 0.07 | 0.07 | 0.07 | 0.06 |  |  |  |  | 22
| 1 | 3 | 0.1 | 0.08 | 0.095272 | 0.08 | 0.2 | 0.09 | 0.11 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.07 | < 5 | 0.07 | < 0.1 | 0.07 | < 0.1 | < 0.1 | < 0.11 | 0.07 | 0.0469 | 0.07 | 0.08 | 0.06 | 0.06 |  |  |  |  | 16
| 1 | 4 | 0.1 | 0.08 | 0.10042 | 0.07 | 0.2 | 0.08 | 0.09 | < 0.1 | 0.07 | 0.07 | 0.08 | 0.07 | < 5 | 0.07 | < 0.1 | 0.08 | 0.28 | < 0.1 | < 0.11 | 0.07 | 0.0448 | 0.07 | 0.06 | 0.07 | 0.06 |  |  |  |  | 0.072995
| 1 | 5 | 0.07 | 0.07 | 0.093578 | 0.07 | 0.3 | 0.09 | 0.08 | < 0.1 | 0.06 | 0.07 | 0.07 | 0.05 | < 5 | 0.06 | < 0.1 | 0.08 | < 0.1 | < 0.1 | < 0.11 | 0.07 | 0.0409 | 0.07 | 0.06 | 0.06 | 0.05 |  |  |  |  | 20
| 1 | 6 | 0.11 | 0.08 | 0.093669 | 0.09 | 0.3 | 0.09 | 0.08 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.06 | < 5 | 0.08 | < 0.1 | 0.07 | < 0.1 | < 0.1 | < 0.11 | 0.07 | 0.0385 | 0.07 | 0.09 | 0.06 | 0.06
| Mean |  | 0.096667 | 0.076667 | 0.094447 | 0.075 | 0.25 | 0.088333 | 0.083333 |  | 0.066667 | 0.07 | 0.073333 | 0.063333 |  | 0.07 |  | 0.075 | 0.28 |  |  | 0.07 | 0.044667 | 0.07 | 0.075 | 0.065 | 0.058333
| Median |  | 0.1 | 0.08 | 0.093623 | 0.07 | 0.25 | 0.09 | 0.08 |  | 0.07 | 0.07 | 0.07 | 0.065 |  | 0.07 |  | 0.075 | 0.28 |  |  | 0.07 | 0.04575 | 0.07 | 0.075 | 0.065 | 0.06
| Std Dev. |  | 0.013663 | 0.005164 | 0.0032365 | 0.0083666 | 0.054772 | 0.0040825 | 0.015055 |  | 0.005164 | 0 | 0.005164 | 0.008165 |  | 0.0089443 |  | 0.0054772 |  |  |  | 0 | 0.0042889 | 0 | 0.013784 | 0.0054772 | 0.0040825
| Rel.Std.Dev. |  | 14.13% | 6.736% | 3.427% | 11.16% | 21.91% | 4.622% | 18.07% |  | 7.746% | 0% | 7.042% | 12.89% |  | 12.78% |  | 7.303% |  |  |  | 0% | 9.602% | 0% | 18.38% | 8.427% | 6.999%
| PDM3 |  | 32.43% | 5.03% | 29.39% | 2.747% | 242.5% | 21.01% | 14.16% |  | -8.67% | -4.103% | 0.4635% | -13.24% |  | -4.103% |  | 2.747% | 283.6% |  |  | -4.103% | -38.81% | -4.103% | 2.747% | -10.95% | -20.09%
| Z-Score (Absolute) |  | 1.96 | 0.49 | 1.8 | 0.37 | NA | 1.35 | 0.98 | 1.47 | 0.25 | 0 | 0.25 | 0.49 | 178.75 | 0 | 1.47 | 0.37 | 1.35 | 1.47 | 1.1 | 0 | 1.86 | 0 | 0.37 | 0.37 | 0.86
Results from laboratory 1.08 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 240b (Certified Value 5.66 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.47 | 5.7734 | 5.771 | 5.731 | 5.26 | 6.311 | 5.83 | 5.6899 | 5.46 | 5.57 | 5.54 | 5.76 | 5.76 | 5.59 | 5.78 | 5.809 | 5.5 | 5.654 | 5.33 | 5.4117 | 5.82 | 5.82 | 6.0507 | 5.48 | 5.33 | 5.49 | 5.87 |  |  | 1
| 1 | 2 | 5.55 | 5.8186 | 5.8031 | 5.7772 | 5.21 | 6.346 | 5.82 | 5.6465 | 5.49 | 5.56 | 5.57 | 6.01 | 5.83 | 5.56 | 5.68 | 5.816 | 5.5 | 5.622 | 4.82 | 5.4798 | 5.801 | 5.96 | 6.1387 | 5.61 | 5.12 | 5.45 | 6.07 |  |  | #N/A
| 1 | 3 | 5.54 | 5.9146 | 5.7455 | 5.8045 | 5.38 | 6.275 | 5.84 | 5.9479 | 5.44 | 5.61 | 5.52 | 6.02 | 5.72 | 5.51 | 5.72 | 5.83 | 5.53 | 5.533 | 5.2 | 5.5426 | 5.761 | 5.95 | 6.1157 | 5.5 | 5.31 | 5.48 | 6.04 |  |  | 16
| 1 | 4 | 5.5 | 5.8741 | 5.7565 | 5.6958 | 5.29 | 6.375 | 5.84 | 5.7083 | 5.41 | 5.62 | 5.47 | 6.03 | 5.74 | 5.62 | 5.67 | 5.832 | 5.42 | 5.534 | 4.62 | 5.6415 | 5.781 | 5.9 | 6.1017 | 5.59 | 5.27 | 5.45 | 5.78 |  |  | 5.6632
| 1 | 5 | 5.39 | 5.678 | 5.7712 | 5.658 | 5.82 | 6.375 | 5.76 | 5.5547 | 5.44 | 5.72 | 5.69 | 6.04 | 5.48 | 5.56 | 5.61 | 5.818 | 5.39 | 5.798 | 5.42 | 5.615 | 5.788 | 5.93 | 6.0087 | 5.62 | 5.14 | 5.49 | 6.2 |  |  | 21
| 1 | 6 | 5.54 | 5.9929 | 5.7932 | 5.7071 | 5.2 | 6.325 | 5.74 | 5.6481 | 5.45 | 5.73 | 5.66 | 5.73 | 5.56 | 5.6 | 5.49 | 5.841 | 5.37 | 5.681 | 4.38 | 5.6164 | 5.831 | 5.98 | 6.0137 | 5.5 | 5.24 | 5.42 | 6.08
| Mean |  | 5.4983 | 5.8419 | 5.7734 | 5.7289 | 5.36 | 6.3345 | 5.805 | 5.6992 | 5.4483 | 5.635 | 5.575 | 5.9317 | 5.6817 | 5.5733 | 5.6583 | 5.8243 | 5.4517 | 5.637 | 4.9617 | 5.5512 | 5.797 | 5.9233 | 6.0715 | 5.55 | 5.235 | 5.4633 | 6.0067
| Median |  | 5.52 | 5.8464 | 5.7711 | 5.7191 | 5.275 | 6.3355 | 5.825 | 5.669 | 5.445 | 5.615 | 5.555 | 6.015 | 5.73 | 5.575 | 5.675 | 5.824 | 5.46 | 5.638 | 5.01 | 5.5788 | 5.7945 | 5.94 | 6.0762 | 5.545 | 5.255 | 5.465 | 6.055
| Std Dev. |  | 0.061128 | 0.11065 | 0.021677 | 0.05413 | 0.23452 | 0.038976 | 0.043704 | 0.13288 | 0.026394 | 0.073417 | 0.084558 | 0.14525 | 0.13303 | 0.038816 | 0.099883 | 0.011944 | 0.066758 | 0.099796 | 0.41897 | 0.090673 | 0.025822 | 0.057504 | 0.054948 | 0.063246 | 0.08735 | 0.028048 | 0.15358
| Rel.Std.Dev. |  | 1.112% | 1.894% | 0.3755% | 0.9448% | 4.375% | 0.6153% | 0.7529% | 2.332% | 0.4844% | 1.303% | 1.517% | 2.449% | 2.341% | 0.6965% | 1.765% | 0.2051% | 1.225% | 1.77% | 8.444% | 1.633% | 0.4454% | 0.9708% | 0.905% | 1.14% | 1.669% | 0.5134% | 2.557%
| PDM3 |  | -2.911% | 3.156% | 1.947% | 1.161% | -5.354% | 11.85% | 2.504% | 0.6366% | -3.794% | -0.4976% | -1.557% | 4.741% | 0.3264% | -1.587% | -0.08559% | 2.846% | -3.735% | -0.4623% | -12.39% | -1.978% | 2.363% | 4.594% | 7.211% | -1.999% | -7.561% | -3.529% | 6.065%
| Z-Score (Absolute) |  | 0.67 | 0.77 | 0.49 | 0.3 | 1.26 | 2.85 | 0.62 | 0.17 | 0.89 | 0.1 | 0.35 | 1.15 | 0.1 | 0.36 | 0 | 0.7 | 0.87 | 0.09 | 2.94 | 0.45 | 0.58 | 1.12 | 1.74 | 0.46 | 1.78 | 0.82 | 1.47
| Analytical results for Cd in OREAS 240b (Certified Value 0.89 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.01 | 0.94 | 0.3 | 0.96883 | 0.7 | 0.9 | 0.98 | 0.81 | 1.3 | 0.8 | 0.85 | 0.9 | 0.81 | 0.94 | < 1 | 0.9 | 0.8 | 0.82 | < 0.1 | 1 | 0.83683 | 0.86 | 0.8761 | 0.9 | 1 | 0.92 | 0.92 |  |  | 1
| 1 | 2 | 1.02 | 0.87 | 0.32 | 0.9964 | 0.7 | 1.1 | 1.06 | 0.91 | 1.3 | 0.9 | 0.87 | 0.91 | 0.85 | 0.85 | < 1 | 0.87 | 0.8 | 0.88 | < 0.1 | 1 | 0.85293 | 0.88 | 0.9141 | 0.85 | 0.97 | 0.9 | 0.92 |  |  | 23
| 1 | 3 | 0.99 | 0.92 | 0.29 | 0.9658 | 0.7 | 1 | 0.86 | 0.89 | 1.3 | 0.9 | 0.82 | 0.92 | 0.82 | 0.86 | < 1 | 0.8 | 0.8 | 0.84 | < 0.1 | 1 | 0.80178 | 0.9 | 0.9078 | 0.85 | 0.98 | 0.9 | 0.86 |  |  | 16
| 1 | 4 | 1.01 | 0.91 | 0.29 | 0.95735 | 0.6 | 1 | 0.98 | 0.84 | 1.3 | 1 | 0.85 | 0.87 | 0.88 | 0.92 | < 1 | 0.8 | 0.8 | 0.83 | < 0.1 | 1 | 0.82351 | 0.89 | 0.9038 | 0.85 | 0.94 | 0.92 | 0.97 |  |  | 0.88591
| 1 | 5 | 0.98 | 0.87 | 0.31 | 0.96681 | 0.7 | 1 | 1.01 | 0.9 | 1.3 | 1 | 0.85 | 0.89 | 0.81 | 0.88 | < 1 | 0.82 | 0.8 | 0.82 | < 0.1 | 1 | 0.82278 | 0.88 | 0.8215 | 0.9 | 0.95 | 0.88 | 0.84 |  |  | 22
| 1 | 6 | 1.04 | 0.94 | 0.3 | 0.96057 | 0.7 | 1.2 | 0.96 | 0.9 | 1.3 | 0.9 | 0.86 | 0.86 | 0.81 | 0.87 | < 1 | 0.87 | 0.8 | 0.84 | < 0.1 | 1 | 0.89065 | 0.86 | 0.8497 | 0.86 | 0.92 | 0.87 | 0.91
| Mean |  | 1.0083 | 0.90833 | 0.30167 | 0.96929 | 0.68333 | 1.0333 | 0.975 | 0.875 | 1.3 | 0.91667 | 0.85 | 0.89167 | 0.83 | 0.88667 |  | 0.84333 | 0.8 | 0.83833 |  | 1 | 0.83808 | 0.87833 | 0.87883 | 0.86833 | 0.96 | 0.89833 | 0.90333
| Median |  | 1.01 | 0.915 | 0.3 | 0.9663 | 0.7 | 1 | 0.98 | 0.895 | 1.3 | 0.9 | 0.85 | 0.895 | 0.815 | 0.875 |  | 0.845 | 0.8 | 0.835 |  | 1 | 0.83017 | 0.88 | 0.88995 | 0.855 | 0.96 | 0.9 | 0.915
| Std Dev. |  | 0.02137 | 0.031885 | 0.01169 | 0.013942 | 0.040825 | 0.10328 | 0.066257 | 0.040373 | 0 | 0.075277 | 0.016733 | 0.023166 | 0.028983 | 0.03559 |  | 0.042269 | 1.2162e-16 | 0.022286 |  | 0 | 0.030821 | 0.016021 | 0.037012 | 0.024833 | 0.028983 | 0.020412 | 0.046762
| Rel.Std.Dev. |  | 2.119% | 3.51% | 3.875% | 1.438% | 5.974% | 9.995% | 6.796% | 4.614% | 0% | 8.212% | 1.969% | 2.598% | 3.492% | 4.014% |  | 5.012% | 1.52e-14% | 2.658% |  | 0% | 3.678% | 1.824% | 4.212% | 2.86% | 3.019% | 2.272% | 5.177%
| PDM3 |  | 13.82% | 2.531% | -65.95% | 9.412% | -22.87% | 16.64% | 10.06% | -1.231% | 46.74% | 3.472% | -4.053% | 0.6499% | -6.311% | 0.08553% |  | -4.806% | -9.697% | -5.37% |  | 12.88% | -5.399% | -0.8551% | -0.7987% | -1.984% | 8.363% | 1.402% | 1.967%
| Z-Score (Absolute) |  | 2.5 | 0.58 | 11.11 | 1.75 | NA | NA | 1.86 | 0.06 | NA | NA | 0.55 | 0.26 | 0.93 | 0.16 | 7.29 | 0.67 | NA | 0.77 | 15.96 | NA | 0.78 | 0 | 0.01 | 0.19 | 1.57 | 0.39 | 0.48
Results from laboratories 1.07, 1.08, 1.12, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 240b (Certified Value 14.9 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 15.31 | 14.79 | 13.43 | 13.8487 | 15.2533 | 13.9 | 17 | 14.26 | 15 | 14.6 | 15.65 | 14.1 | 14.2 | 12 | 14 | 14.17 | 16.5 | 14.617 | 15.7 | 8.8978 | 17.8 | 15.73 | 15.51 | 17.5 | [5 empty] | 1
| 1 | 2 | 15.01 | 14.65 | 14.17 | 13.655 | 15.4401 | 14.2 | 17 | 14.3 | 15 | 14.95 | 16.1 | 15.85 | 14.3 | < 10 | 14 | 14.29 | 16.5 | 14.3786 | 15.35 | 9.1 | 17.4 | 15.24 | 15.45 | 16.7 | [5 empty] | 24
| 1 | 3 | 14.79 | 15.15 | 14.08 | 13.7962 | 15.5673 | 14.1 | 17 | 14.55 | 15 | 14.35 | 15.6 | 14.35 | 14.65 | 12 | 14 | 14.7 | 16 | 15.1288 | 16.05 | 9.0644 | 17.5 | 15.17 | 16.22 | 15.7 | [5 empty] | 16
| 1 | 4 | 14.72 | 14.8 | 13.28 | 13.8757 | 15.5236 | 14.3 | 17 | 14.25 | 15 | 14.45 | 15.25 | 14.3 | 15 | < 10 | 14 | 13.89 | 15.5 | 15.0727 | 15.55 | 8.8745 | 17.3 | 14.91 | 15.92 | 17.8 | [5 empty] | 14.8792
| 1 | 5 | 13.59 | 14.42 | 13.44 | 13.4266 | 15.649 | 15 | 17 | 13.98 | 15 | 14.4 | 15.45 | 14 | 14.2 | < 10 | 14 | 14.43 | 15.5 | 14.9067 | 15.35 | 8.4164 | 17.8 | 14.58 | 15.89 | 15.3 | [5 empty] | 23
| 1 | 6 | 14.42 | 15.41 | 14.06 | 13.7615 | 15.5321 | 14 | 17 | 14.09 | 15 | 15.05 | 15.35 | 14.15 | 13.95 | < 10 | 14 | 14.73 | 16 | 15.7318 | 15.4 | 7.9373 | 17.9 | 14.73 | 16.18 | 16.9
| Mean |  | 14.64 | 14.87 | 13.7433 | 13.7273 | 15.4942 | 14.25 | 17 | 14.2383 | 15 | 14.6333 | 15.5667 | 14.4583 | 14.3833 | 12 | 14 | 14.3683 | 16 | 14.9726 | 15.5667 | 8.7151 | 17.6167 | 15.06 | 15.8617 | 16.65
| Median |  | 14.755 | 14.795 | 13.75 | 13.7788 | 15.5278 | 14.15 | 17 | 14.255 | 15 | 14.525 | 15.525 | 14.225 | 14.25 | 12 | 14 | 14.36 | 16 | 14.9897 | 15.475 | 8.8862 | 17.65 | 15.04 | 15.905 | 16.8
| Std Dev. |  | 0.59424 | 0.35558 | 0.40013 | 0.16633 | 0.13605 | 0.3937 | 0 | 0.19508 | 0 | 0.29777 | 0.30111 | 0.69384 | 0.37771 | 0 | 0 | 0.32214 | 0.44721 | 0.46785 | 0.27325 | 0.45236 | 0.24833 | 0.41357 | 0.32468 | 0.98336
| Rel.Std.Dev. |  | 4.059% | 2.391% | 2.911% | 1.212% | 0.8781% | 2.763% | 0% | 1.37% | 0% | 2.035% | 1.934% | 4.799% | 2.626% | 0% | 0% | 2.242% | 2.795% | 3.125% | 1.755% | 5.191% | 1.41% | 2.746% | 2.047% | 5.906%
| PDM3 |  | -1.608% | -0.06191% | -7.634% | -7.742% | 4.133% | -4.229% | 14.25% | -4.307% | 0.8118% | -1.652% | 4.62% | -2.829% | -3.333% | -19.35% | -5.909% | -3.434% | 7.533% | 0.6277% | 4.62% | -41.43% | 18.4% | 1.215% | 6.603% | 11.9%
| Z-Score (Absolute) |  | 0.12 | 0.12 | 1.09 | 1.1 | 0.79 | 0.54 | NA | 0.55 | NA | 0.13 | 0.87 | 0.32 | 0.4 | NA | NA | 0.42 | 1.34 | 0.23 | 0.87 | 6.49 | 3.07 | 0.33 | 1.19 | 2.03
Results from laboratories 1.08, 1.13, 1.19 and 1.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 240b (Certified Value 40.3 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 43.5 | 40.3 | 38.44 | 33.361 | 40.2674 | 40 | 42.5 | 41 | 39.7 | 31 | 39.8 | 40.9 | 42.9 | 40 | 39.9 | 41 | 37.7 | 39 | 42.9 | 33.96 | 35 | 41.2054 | 40.6 | 35.8266 | 46.6 | 41.8 | 37.4 | 43.1 |  | 1
| 1 | 2 | 44.3 | 40.2 | 39.91 | 32.5586 | 40.7439 | 40 | 48.5 | 40 | 40 | 31 | 39.4 | 41.9 | 43.7 | 42.7 | 39.8 | 40 | 41 | 38.3 | 41.6 | 33.28 | 35 | 41.8197 | 40.7 | 38.9881 | 46.4 | 41.8 | 37 | 41.7 |  | 25
| 1 | 3 | 43.5 | 41 | 39.09 | 32.4092 | 40.7381 | 40 | 45.4 | 41 | 41.2 | 31 | 38.3 | 40.8 | 43.1 | 40.1 | 41.7 | 40 | 39.7 | 38.6 | 40.9 | 32.02 | 35 | 40.8023 | 41.6 | 37.6141 | 46.6 | 43.6 | 37.7 | 38.9 |  | 16
| 1 | 4 | 44 | 40.5 | 39.62 | 32.8651 | 40.0955 | 40 | 46.9 | 42 | 40.2 | 31 | 39.2 | 40.2 | 41.7 | 40.6 | 42 | 40 | 41.1 | 37.5 | 40.5 | 34.16 | 35 | 41.4198 | 40.9 | 37.0617 | 46.3 | 41.1 | 38.3 | 43.9 |  | 40.2925
| 1 | 5 | 42 | 40.8 | 38.82 | 32.0702 | 39.8298 | 40 | 46.2 | 41 | 39.2 | 31 | 40.4 | 40.2 | 42.5 | 39.2 | 39.6 | 40 | 38.1 | 36.8 | 40.5 | 36.55 | 35 | 41.2114 | 41.5 | 36.1645 | 47.7 | 40.9 | 37.2 | 37.9 |  | 24
| 1 | 6 | 42.5 | 41.1 | 39.66 | 32.3985 | 39.9456 | 39 | 52.3 | 40 | 39.8 | 31 | 39.5 | 41.5 | 43.6 | 39.5 | 41 | 39 | 39.8 | 37.3 | 41.3 | 32.76 | 35 | 42.2857 | 40.3 | 35.0288 | 47.4 | 41.1 | 37.6 | 42.3
| Mean |  | 43.3 | 40.65 | 39.2567 | 32.6105 | 40.2701 | 39.8333 | 46.9667 | 40.8333 | 40.0167 | 31 | 39.4333 | 40.9167 | 42.9167 | 40.35 | 40.6667 | 40 | 39.5667 | 37.9167 | 41.2833 | 33.7883 | 35 | 41.4574 | 40.9333 | 36.7806 | 46.8333 | 41.7167 | 37.5333 | 41.3
| Median |  | 43.5 | 40.65 | 39.355 | 32.4839 | 40.1815 | 40 | 46.55 | 41 | 39.9 | 31 | 39.45 | 40.85 | 43 | 40.05 | 40.45 | 40 | 39.75 | 37.9 | 41.1 | 33.62 | 35 | 41.3156 | 40.8 | 36.6131 | 46.6 | 41.45 | 37.5 | 42
| Std Dev. |  | 0.88318 | 0.37283 | 0.56691 | 0.44902 | 0.39324 | 0.40825 | 3.28 | 0.75277 | 0.67057 | 0 | 0.69474 | 0.68532 | 0.74409 | 1.2502 | 1.0424 | 0.63246 | 1.4222 | 0.84715 | 0.90425 | 1.5644 | 0 | 0.52381 | 0.5164 | 1.4146 | 0.57504 | 0.99883 | 0.45461 | 2.3866
| Rel.Std.Dev. |  | 2.04% | 0.9172% | 1.444% | 1.377% | 0.9765% | 1.025% | 6.984% | 1.844% | 1.676% | 0% | 1.762% | 1.675% | 1.734% | 3.098% | 2.563% | 1.581% | 3.594% | 2.234% | 2.19% | 4.63% | 0% | 1.263% | 1.262% | 3.846% | 1.228% | 2.394% | 1.211% | 5.779%
| PDM3 |  | 7.464% | 0.8872% | -2.571% | -19.07% | -0.05567% | -1.14% | 16.56% | 1.342% | -0.6846% | -23.06% | -2.132% | 1.549% | 6.513% | 0.1427% | 0.9286% | -0.726% | -1.801% | -5.896% | 2.459% | -16.14% | -13.14% | 2.891% | 1.59% | -8.716% | 16.23% | 3.535% | -6.848% | 2.5%
| Z-Score (Absolute) |  | 1.97 | 0.22 | 0.7 | 5.08 | 0.03 | 0.31 | 4.39 | 0.35 | 0.19 | 6.14 | 0.58 | 0.4 | 1.72 | 0.03 | 0.24 | 0.2 | 0.49 | 1.58 | 0.64 | 4.3 | 3.5 | 0.76 | 0.41 | 2.33 | 4.31 | 0.93 | 1.83 | 0.65
| Analytical results for Cr in OREAS 240b (Certified Value 146 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 118 | 176 | 138.2315 | 145 | 205 | 130 | 123 | 152 | 163 | 132 | 146 | 141 | 132 | 151 | 120 | 144 | 160 | 141.23 | 141.651 | 165 | 171.0356 | 137 | 149.6968 | 151 | 155 | 160 | 145 |  |  | 1
| 1 | 2 | 125 | 172 | 138.4937 | 92 | 205 | 130 | 120 | 150 | 160 | 137 | 149 | 136 | 131 | 145 | 131 | 144 | 143 | 135.52 | 156.663 | 165 | 171.1318 | 142 | 153.1935 | 148 | 150 | 155 | 154 |  |  | 26
| 1 | 3 | 119 | 172 | 134.7765 | 96 | 206 | 140 | 121 | 151 | 157 | 136 | 150 | 141 | 129 | 155 | 128 | 145 | 148 | 143.47 | 155.799 | 160 | 168.7902 | 141 | 151.7968 | 143 | 154 | 157 | 141 |  |  | 16
| 1 | 4 | 122 | 173 | 134.2745 | 116 | 209 | 132 | 139 | 150 | 160 | 132 | 153 | 139 | 130 | 150 | 116 | 142 | 151 | 137.24 | 160.146 | 165 | 174.137 | 138 | 151.4976 | 145 | 152 | 157 | 129 |  |  | 145.6833
| 1 | 5 | 112 | 176 | 135.3951 | 112 | 207 | 150 | 116 | 152 | 170 | 135 | 153 | 132 | 130 | 142 | 122 | 141 | 168 | 149.87 | 158.913 | 150 | 173.1354 | 143 | 149.505 | 136 | 150 | 162 | 144 |  |  | 25
| 1 | 6 | 115 | 175 | 133.6944 | 98 | 209 | 138 | 122 | 152 | 174 | 135 | 137 | 133 | 129 | 143 | 125 | 143 | 149 | 140.23 | 160.101 | 165 | 178.6913 | 144 | 149.1065 | 144 | 149 | 161 | 135
| Mean |  | 118.5 | 174 | 135.811 | 109.8333 | 206.8333 | 136.6667 | 123.5 | 151.1667 | 164 | 134.5 | 148 | 137 | 130.1667 | 147.6667 | 123.6667 | 143.1667 | 153.1667 | 141.26 | 155.5455 | 161.6667 | 172.8202 | 140.8333 | 150.7994 | 144.5 | 151.6667 | 158.6667 | 141.3333
| Median |  | 118.5 | 174 | 135.0858 | 105 | 206.5 | 135 | 121.5 | 151.5 | 161.5 | 135 | 149.5 | 137.5 | 130 | 147.5 | 123.5 | 143.5 | 150 | 140.73 | 157.788 | 165 | 172.1336 | 141.5 | 150.5972 | 144.5 | 151 | 158.5 | 142.5
| Std Dev. |  | 4.6797 | 1.8974 | 2.0562 | 19.6409 | 1.8348 | 7.7632 | 7.9687 | 0.98319 | 6.603 | 2.0736 | 6 | 3.9497 | 1.169 | 5.1251 | 5.465 | 1.472 | 9.1524 | 5.083 | 7.0374 | 6.0553 | 3.4229 | 2.7869 | 1.6106 | 5.0892 | 2.4221 | 2.7325 | 8.641
| Rel.Std.Dev. |  | 3.949% | 1.09% | 1.514% | 17.88% | 0.8871% | 5.68% | 6.452% | 0.6504% | 4.026% | 1.542% | 4.054% | 2.883% | 0.8981% | 3.471% | 4.419% | 1.028% | 5.975% | 3.598% | 4.524% | 3.746% | 1.981% | 1.979% | 1.068% | 3.522% | 1.597% | 1.722% | 6.114%
| PDM3 |  | -18.66% | 19.44% | -6.777% | -24.61% | 41.97% | -6.189% | -15.23% | 3.764% | 12.57% | -7.676% | 1.59% | -5.96% | -10.65% | 1.361% | -15.11% | -1.727% | 5.137% | -3.036% | 6.77% | 10.97% | 18.63% | -3.329% | 3.512% | -0.8122% | 4.107% | 8.912% | -2.986%
| Z-Score (Absolute) |  | 2.02 | 2.29 | 0.67 | 2.69 | 4.84 | 0.61 | 1.63 | 0.52 | 1.51 | 0.78 | 0.27 | 0.58 | 1.11 | 0.25 | 1.62 | 0.1 | 0.67 | 0.25 | 0.86 | 1.33 | 2.2 | 0.28 | 0.49 | 0 | 0.56 | 1.1 | 0.25
| Analytical results for Cs in OREAS 240b (Certified Value 1.38 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | < 5 | 1.34 | 1.3016 | 1.24 | 1.4 | 1.3 | 1.35 | 1.4 | 1.33 | 1.39 | 1.36 | 1.37 | 1.3 | 1.37 | 1.5 | 1.256 | 1.45 | 1.3933 | 1.31 | 1.59 | 1.41 | 1.47 | [7 empty] | 1
| 1 | 2 | < 5 | 1.39 | 1.2724 | 1.25 | 1.5 | 1.3 | 1.4 | 1.5 | 1.34 | 1.44 | 1.44 | 1.29 | 1.5 | 1.33 | 1.5 | 1.193 | 1.38 | 1.3968 | 1.31 | 1.5 | 1.36 | 1.42 | [7 empty] | 27
| 1 | 3 | < 5 | 1.37 | 1.3051 | 1.29 | 1.4 | 1.3 | 1.47 | 1.4 | 1.28 | 1.4 | 1.34 | 1.33 | 1.4 | 1.32 | 1.5 | 1.3234 | 1.48 | 1.4085 | 1.31 | 1.58 | 1.42 | 1.36 | [7 empty] | 16
| 1 | 4 | < 5 | 1.39 | 1.3339 | 1.26 | 1.4 | 1.3 | 1.36 | 1.5 | 1.33 | 1.34 | 1.36 | 1.35 | 1.4 | 1.35 | 1.5 | 1.2811 | 1.46 | 1.3886 | 1.31 | 1.47 | 1.44 | 1.52 | [7 empty] | 1.3785
| 1 | 5 | < 5 | 1.37 | 1.2168 | 1.4 | 1.4 | 1.3 | 1.35 | 1.6 | 1.34 | 1.38 | 1.38 | 1.3 | 1.3 | 1.29 | 1.5 | 1.2713 | 1.43 | 1.3197 | 1.31 | 1.53 | 1.41 | 1.3 | [7 empty] | 26
| 1 | 6 | < 5 | 1.37 | 1.2948 | 1.27 | 1.6 | 1.3 | 1.36 | 1.5 | 1.35 | 1.4 | 1.32 | 1.33 | 1.5 | 1.31 | 1.5 | 1.369 | 1.44 | 1.257 | 1.3 | 1.51 | 1.4 | 1.42
| Mean |  |  | 1.3717 | 1.2874 | 1.285 | 1.45 | 1.3 | 1.3817 | 1.4833 | 1.3283 | 1.3917 | 1.3667 | 1.3283 | 1.4 | 1.3283 | 1.5 | 1.2823 | 1.44 | 1.3606 | 1.3083 | 1.53 | 1.4067 | 1.415
| Median |  |  | 1.37 | 1.2982 | 1.265 | 1.4 | 1.3 | 1.36 | 1.5 | 1.335 | 1.395 | 1.36 | 1.33 | 1.4 | 1.325 | 1.5 | 1.2762 | 1.445 | 1.391 | 1.31 | 1.52 | 1.41 | 1.42
| Std Dev. |  |  | 0.018348 | 0.039851 | 0.058907 | 0.083666 | 0 | 0.047081 | 0.075277 | 0.024833 | 0.032506 | 0.041312 | 0.029944 | 0.089443 | 0.028577 | 0 | 0.059996 | 0.034059 | 0.059773 | 0.0040825 | 0.046904 | 0.026583 | 0.07791
| Rel.Std.Dev. |  |  | 1.338% | 3.095% | 4.584% | 5.77% | 0% | 3.408% | 5.075% | 1.869% | 2.336% | 3.023% | 2.254% | 6.389% | 2.151% | 0% | 4.679% | 2.365% | 4.393% | 0.312% | 3.066% | 1.89% | 5.506%
| PDM3 |  |  | -0.4982% | -6.609% | -6.785% | 5.184% | -5.697% | 0.2272% | 7.602% | -3.642% | 0.9526% | -0.8609% | -3.642% | 1.557% | -3.642% | 8.811% | -6.981% | 4.459% | -1.297% | -5.092% | 10.99% | 2.041% | 2.645%
| Z-Score (Absolute) |  | 13.57 | 0.06 | 1.08 | 1.11 | 0.89 | 0.93 | 0.06 | 1.29 | 0.58 | 0.18 | 0.12 | 0.58 | 0.28 | 0.58 | 1.49 | 1.14 | 0.76 | 0.19 | 0.83 | 1.85 | 0.36 | 0.46
| Analytical results for Cu in OREAS 240b (Certified Value 164 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 167.4 | 163.7 | 151.49 | 177.79 | 165.8365 | 168 | 163 | 161 | 158.6 | 160 | 161.7 | 163.5 | 164.5 | 165.5 | 168 | 179 | 163 | 151 | 170 | 27.09 | 158.571 | 160 | 155.3645 | 161.5 | 120.417 | 177 | 168.4 | 164.1 | 154 | 1
| 1 | 2 | 172.4 | 163 | 154.27 | 178.98 | 167.427 | 168 | 183 | 161 | 158.3 | 160 | 163.7 | 166.5 | 173 | 169.5 | 165 | 171 | 167 | 147.7 | 168 | 29.39 | 158.13 | 160 | 158.8875 | 162 | 128.1096 | 179 | 163.2 | 164.3 | 157 | 28
| 1 | 3 | 172.7 | 162.1 | 154.41 | 180.4 | 166.4863 | 167 | 174 | 163 | 167 | 161 | 169.2 | 162 | 170.5 | 163.5 | 168 | 172 | 167 | 150.3 | 162 | 22.64 | 162.567 | 160 | 157.4126 | 166 | 122.5737 | 179 | 169.6 | 166.4 | 157 | 16
| 1 | 4 | 170 |  | 158.77 | 179.84 | 164.9751 | 173 | 180 | 163 | 161.8 | 161 | 169.5 | 161.5 | 169.5 | 166 | 169 | 172 | 165 | 148.6 | 167 | 30.55 | 163.629 | 155 | 157.5135 | 165.5 | 120.1402 | 176 | 162.7 | 161.9 | 152 | 164.3747
| 1 | 5 | 165.1 | 163.7 | 152.21 | 175.99 | 163.7908 | 166 | 180 | 155 | 153.6 | 162 | 167.6 | 166 | 171 | 160.5 | 167 | 171 | 163 | 145.7 | 173 | 34.67 | 168.201 | 155 | 159.8958 | 165 | 114.4309 | 177 | 161.9 | 166.1 | 160 | 27
| 1 | 6 | 173.6 | 163.2 | 157.2 | 178.98 | 164.4558 | 165 | 204 | 168 | 155.4 | 160 | 167.1 | 165 | 167 | 160 | 166.5 | 167 | 166 | 148.2 | 168 | 30.03 | 166.176 | 160 | 161.6454 | 166 | 109.8731 | 177 | 166.4 | 163 | 155
| Mean |  | 170.2 | 163.14 | 154.725 | 178.6633 | 165.4952 | 167.8333 | 180.6667 | 161.8333 | 159.1167 | 160.6667 | 166.4667 | 164.0833 | 169.25 | 164.1667 | 167.25 | 172 | 165.1667 | 148.5833 | 168 | 29.0617 | 162.879 | 158.3333 | 158.4532 | 164.3333 | 119.2574 | 177.5 | 165.3667 | 164.3 | 155.8333
| Median |  | 171.2 | 163.2 | 154.34 | 178.98 | 165.4058 | 167.5 | 180 | 162 | 158.45 | 160.5 | 167.35 | 164.25 | 170 | 164.5 | 167.5 | 171.5 | 165.5 | 148.4 | 168 | 29.71 | 163.098 | 160 | 158.2005 | 165.25 | 120.2786 | 177 | 164.8 | 164.2 | 156
| Std Dev. |  | 3.3627 | 0.65803 | 2.8137 | 1.5818 | 1.3481 | 2.7869 | 13.471 | 4.2151 | 4.786 | 0.8165 | 3.122 | 2.0837 | 3.0455 | 3.6009 | 1.4053 | 3.8987 | 1.8348 | 1.8989 | 3.6332 | 3.996 | 4.0241 | 2.582 | 2.1893 | 2.0412 | 6.3694 | 1.2247 | 3.2253 | 1.7401 | 2.7869
| Rel.Std.Dev. |  | 1.976% | 0.4034% | 1.819% | 0.8854% | 0.8146% | 1.661% | 7.456% | 2.605% | 3.008% | 0.5082% | 1.875% | 1.27% | 1.799% | 2.193% | 0.8403% | 2.267% | 1.111% | 1.278% | 2.163% | 13.75% | 2.471% | 1.631% | 1.382% | 1.242% | 5.341% | 0.69% | 1.95% | 1.059% | 1.788%
| PDM3 |  | 3.544% | -0.7511% | -5.871% | 8.693% | 0.6817% | 2.104% | 9.911% | -1.546% | -3.199% | -2.256% | 1.273% | -0.1772% | 2.966% | -0.1266% | 1.749% | 4.639% | 0.4818% | -9.607% | 2.206% | -82.32% | -0.9099% | -3.675% | -3.602% | -0.02516% | -27.45% | 7.985% | 0.6035% | -0.04544% | -5.196%
| Z-Score (Absolute) |  | 1.08 | 0.21 | 1.75 | 2.62 | 0.22 | 0.64 | 2.98 | 0.45 | 0.94 | 0.66 | 0.39 | 0.04 | 0.9 | 0.02 | 0.54 | 1.4 | 0.16 | 2.86 | 0.67 | 24.65 | 0.26 | 1.09 | 1.07 | 0.01 | 8.21 | 2.41 | 0.19 | 0 | 1.54
| Analytical results for Dy in OREAS 240b (Certified Value 3.35 ppm) | [30 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.32 | 3.01 | 3.0757 | 3.3 | 3.12 | 3.24 | 3.3 | 3.3124 | [21 empty] | 1
| 1 | 2 | 3.35 | 3.21 | 3.0872 | 3.4 | 3.52 | 3.31 | 3.5 | 3.282 | [21 empty] | 29
| 1 | 3 | 3.34 | 3.23 | 3.0814 | 3.3 | 3.29 | 3.27 | 3.3 | 3.5905 | [21 empty] | 16
| 1 | 4 | 3.28 | 2.98 | 3.1153 | 3.3 | 3.37 | 3.35 | 3.1 | 3.4855 | [21 empty] | 3.3548
| 1 | 5 | 3.28 | 3 | 3.0832 | 3.5 | 3.47 | 3.37 | 3.4 | 3.3981 | [21 empty] | 28
| 1 | 6 | 3.36 | 3.23 | 3.0687 | 3.3 | 3.88 | 3.36 | 3.5 | 3.5492
| Mean |  | 3.3217 | 3.11 | 3.0852 | 3.35 | 3.4417 | 3.3167 | 3.35 | 3.4363
| Median |  | 3.33 | 3.11 | 3.0823 | 3.3 | 3.42 | 3.33 | 3.35 | 3.4418
| Std Dev. |  | 0.034881 | 0.12474 | 0.016053 | 0.083666 | 0.25717 | 0.052789 | 0.15166 | 0.12614
| Rel.Std.Dev. |  | 1.05% | 4.011% | 0.5203% | 2.497% | 7.472% | 1.592% | 4.527% | 3.671%
| PDM3 |  | -0.9866% | -7.296% | -8.034% | -0.142% | 2.59% | -1.136% | -0.142% | 2.429%
| Z-Score (Absolute) |  | 0.16 | 2.55 | 2.83 | 0.16 | 1.19 | 0.22 | 0.16 | 1.13
| Analytical results for Er in OREAS 240b (Certified Value 2.03 ppm) | [30 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.01 | 1.77 | 1.9172 | 2 | 1.92 | 1.97 | 2.2 | 2.0965 | [21 empty] | 1
| 1 | 2 | 2.01 | 1.9 | 1.9494 | 2 | 2.15 | 1.93 | 2.1 | 2.0955 | [21 empty] | 30
| 1 | 3 | 2.11 | 1.92 | 1.9195 | 2 | 2.03 | 2.08 | 2.1 | 2.2422 | [21 empty] | 16
| 1 | 4 | 2.06 | 1.79 | 1.9722 | 2 | 2.06 | 2.02 | 1.9 | 2.1881 | [21 empty] | 2.0298
| 1 | 5 | 2.01 | 1.78 | 1.8905 | 2.1 | 2.08 | 2.02 | 2.1 | 2.186 | [21 empty] | 29
| 1 | 6 | 2.12 | 1.94 | 1.9679 | 1.9 | 2.35 | 2.03 | 2.2 | 2.3466
| Mean |  | 2.0533 | 1.85 | 1.9361 | 2 | 2.0983 | 2.0083 | 2.1 | 2.1925
| Median |  | 2.035 | 1.845 | 1.9345 | 2 | 2.07 | 2.02 | 2.1 | 2.1871
| Std Dev. |  | 0.05164 | 0.077974 | 0.032272 | 0.063246 | 0.14442 | 0.051929 | 0.10954 | 0.094764
| Rel.Std.Dev. |  | 2.515% | 4.215% | 1.667% | 3.162% | 6.883% | 2.586% | 5.216% | 4.322%
| PDM3 |  | 1.158% | -8.859% | -4.615% | -1.47% | 3.375% | -1.059% | 3.457% | 8.014%
| Z-Score (Absolute) |  | 0.22 | 1.78 | 0.93 | 0.3 | 0.67 | 0.22 | 0.68 | 1.6
| Analytical results for Eu in OREAS 240b (Certified Value 0.87 ppm) | [30 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.86 | 0.8 | 1.2717 | 0.81 | 0.82 | 0.92 | 0.9 | 0.86786 | [21 empty] | 1
| 1 | 2 | 0.82 | 0.86 | 1.2876 | 0.81 | 0.94 | 0.9 | 0.9 | 0.8414 | [21 empty] | 31
| 1 | 3 | 0.85 | 0.85 | 1.3437 | 0.81 | 0.88 | 0.94 | 0.8 | 0.88492 | [21 empty] | 16
| 1 | 4 | 0.86 | 0.8 | 1.2958 | 0.83 | 0.88 | 0.94 | 0.9 | 0.86531 | [21 empty] | 0.8694
| 1 | 5 | 0.83 | 0.81 | 1.2416 | 0.92 | 0.91 | 0.96 | 0.9 | 0.89416 | [21 empty] | 30
| 1 | 6 | 0.86 | 0.85 | 1.2642 | 0.83 | 1.02 | 0.96 | 0.9 | 0.91677
| Mean |  | 0.84667 | 0.82833 | 1.2841 | 0.835 | 0.90833 | 0.93667 | 0.88333 | 0.8784
| Median |  | 0.855 | 0.83 | 1.2797 | 0.82 | 0.895 | 0.94 | 0.9 | 0.87639
| Std Dev. |  | 0.017512 | 0.027869 | 0.034812 | 0.042778 | 0.067651 | 0.023381 | 0.040825 | 0.026133
| Rel.Std.Dev. |  | 2.068% | 3.364% | 2.711% | 5.123% | 7.448% | 2.496% | 4.622% | 2.975%
| PDM3 |  | -2.615% | -4.724% | 47.7% | -3.957% | 4.478% | 7.737% | 1.603% | 1.036%
| Z-Score (Absolute) |  | 0.49 | 0.78 | 6.3 | 0.67 | 0.46 | 0.91 | NA | 0
| Analytical results for Fe in OREAS 240b (Certified Value 7.31 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 7.46 | 7.32 | 7.6052 | 7.3541 | 6.74 | 7.974 | 7.6 | 7.15 | 7.13 | 7.08 | 7.39 | 7.25 | 7.35 | 7.3 | 7.21 | 7.62 | 6.83 | 7.28 | 7.51 | 6.949 | 7.485 | 7.3 | 7.3893 | 7.16 | 6.98 | 7.75 | 8.1 |  |  | 1
| 1 | 2 | 7.59 | 7.35 | 7.6066 | 7.4615 | 6.59 | 7.974 | 7.64 | 7.08 | 7.09 | 7.1 | 7.42 | 7.6 | 7.53 | 7.26 | 7.01 | 7.75 | 6.85 | 7.19 | 6.64 | 6.9616 | 7.42 | 7.47 | 7.4673 | 7.33 | 6.96 | 7.73 | 7.6 |  |  | #N/A
| 1 | 3 | 7.61 | 7.52 | 7.6232 | 7.4464 | 6.83 | 7.974 | 7.63 | 7.41 | 7.11 | 7 | 7.33 | 7.52 | 7.31 | 7.27 | 7.07 | 7.77 | 6.86 | 7.03 | 7.34 | 7.0886 | 7.365 | 7.44 | 7.4663 | 7.15 | 7.02 | 7.8 | 6.5 |  |  | 16
| 1 | 4 | 7.55 | 7.37 | 7.6396 | 7.3528 | 6.82 | 8.114 | 7.66 | 7.2 | 7.1 | 7.06 | 7.31 | 7.51 | 7.36 | 7.39 | 7.05 | 7.76 | 6.78 | 7.26 | 6.74 | 7.1689 | 7.41 | 7.45 | 7.4593 | 7.31 | 6.96 | 7.75 | 7.8 |  |  | 7.3127
| 1 | 5 | 7.33 | 7.09 | 7.6329 | 7.3099 | 7.28 | 7.974 | 7.52 | 7.03 | 7.13 | 7.18 | 7.54 | 7.5 | 7.09 | 7.38 | 7 | 7.64 | 6.7 | 7.46 | 7.62 | 7.1938 | 7.43 | 7.41 | 7.3533 | 7.34 | 6.74 | 7.73 | 6.4 |  |  | 31
| 1 | 6 | 7.52 | 7.6 | 7.6139 | 7.3153 | 6.65 | 8.114 | 7.46 | 7.09 | 7.08 | 7.22 | 7.57 | 7.24 | 7.16 | 7.33 | 6.81 | 7.77 | 6.68 | 7.26 | 6.33 | 7.2101 | 7.475 | 7.46 | 7.3303 | 7.19 | 6.94 | 7.71 | 7.4
| Mean |  | 7.51 | 7.375 | 7.6202 | 7.3733 | 6.8183 | 8.0207 | 7.585 | 7.16 | 7.1067 | 7.1067 | 7.4267 | 7.4367 | 7.3 | 7.3217 | 7.025 | 7.7183 | 6.7833 | 7.2467 | 7.03 | 7.0953 | 7.4308 | 7.4217 | 7.411 | 7.2467 | 6.9333 | 7.745 | 7.3
| Median |  | 7.535 | 7.36 | 7.6185 | 7.3535 | 6.78 | 7.974 | 7.615 | 7.12 | 7.105 | 7.09 | 7.405 | 7.505 | 7.33 | 7.315 | 7.03 | 7.755 | 6.805 | 7.26 | 7.04 | 7.1288 | 7.425 | 7.445 | 7.4243 | 7.25 | 6.96 | 7.74 | 7.5
| Std Dev. |  | 0.10296 | 0.17695 | 0.014115 | 0.065264 | 0.2449 | 0.072296 | 0.078422 | 0.13594 | 0.020656 | 0.080664 | 0.10746 | 0.15267 | 0.15672 | 0.054924 | 0.12958 | 0.069113 | 0.077632 | 0.13938 | 0.5293 | 0.11629 | 0.044206 | 0.063061 | 0.061434 | 0.089144 | 0.098522 | 0.030822 | 0.69857
| Rel.Std.Dev. |  | 1.371% | 2.399% | 0.1852% | 0.8851% | 3.592% | 0.9014% | 1.034% | 1.899% | 0.2907% | 1.135% | 1.447% | 2.053% | 2.147% | 0.7502% | 1.845% | 0.8954% | 1.144% | 1.923% | 7.529% | 1.639% | 0.5949% | 0.8497% | 0.829% | 1.23% | 1.421% | 0.398% | 9.569%
| PDM3 |  | 2.697% | 0.8514% | 4.205% | 0.8285% | -6.761% | 9.681% | 3.723% | -2.089% | -2.818% | -2.818% | 1.558% | 1.695% | -0.1742% | 0.122% | -3.935% | 5.546% | -7.24% | -0.9036% | -3.866% | -2.973% | 1.615% | 1.49% | 1.343% | -0.9036% | -5.188% | 5.911% | -0.1742%
| Z-Score (Absolute) |  | 0.67 | 0.19 | 1.07 | 0.18 | 1.8 | 2.5 | 0.94 | 0.58 | 0.77 | 0.77 | 0.38 | 0.41 | 0.08 | 0 | 1.06 | 1.42 | 1.93 | 0.27 | 1.04 | 0.81 | 0.39 | 0.36 | 0.32 | 0.27 | 1.39 | 1.52 | 0.08
| Analytical results for Ga in OREAS 240b (Certified Value 14.8 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 16.4 | 14.78 | 14.1717 | 13.1979 | 12.2 | 13.2 | 14.6 | 14.75 | 18 | 14 | 14.25 | 15.75 | 14.35 | 14.75 | 11 | 13.4 | 14.46 | 13.28 | 12 | 11.6113 | 15.35 | 15.2262 | 18.1 | 16 | 14.42 | 15.9 |  |  |  | 1
| 1 | 2 | 16.9 | 14.73 | 13.8575 | 13.159 | 11.8 | 15 | 15 | 15.04 | 18 | 14 | 14.5 | 16.1 | 14.9 | 14.7 | 11 | 15.1 | 14.88 | 8.98 | 12 | 11.2106 | 15.35 | 16.6216 | 17.7 | 15.2 | 13.79 | 15.3 |  |  |  | 32
| 1 | 3 | 16.6 | 14.82 | 13.876 | 13.1472 | 12.6 | 14.1 | 15.2 | 15.46 | 17 | 14 | 14.05 | 15.55 | 14.7 | 15.25 | 11 | 14.5 | 14.92 | 9.76 | 11 | 13.022 | 15.7 | 16.163 | 17.6 | 15.9 | 14.1 | 14.3 |  |  |  | 16
| 1 | 4 | 16.5 | 14.91 | 13.9649 | 13.3939 | 12.3 | 14.5 | 14.8 | 14.82 | 17 | 14 | 14.05 | 15.1 | 14.55 | 14.85 | 11 | 14.6 | 14.7 | 12.52 | 11 | 11.927 | 15.6 | 15.9164 | 17.6 | 15 | 14.66 | 16.3 |  |  |  | 14.8328
| 1 | 5 | 15.4 | 14.78 | 13.5227 | 12.9398 | 13.3 | 14.5 | 14.9 | 14.58 | 18 | 15 | 14.15 | 15.4 | 14.5 | 14.6 | 10 | 13.7 | 14.81 | 14.39 | 11 | 12.1213 | 15.3 | 15.6924 | 18.3 | 14.8 | 14.28 | 14.1 |  |  |  | 32
| 1 | 6 | 16 | 14.84 | 13.7616 | 13.3678 | 12 | 16.1 | 15.4 | 15.05 | 18 | 14 | 14.2 | 14.05 | 14.3 | 14.45 | 10 | 15.7 | 15.29 | 12.58 | 11 | 12.7426 | 15.35 | 14.9835 | 18.2 | 15.7 | 14.41 | 15.4
| Mean |  | 16.3 | 14.81 | 13.8591 | 13.2009 | 12.3667 | 14.5667 | 14.9833 | 14.95 | 17.6667 | 14.1667 | 14.2 | 15.325 | 14.55 | 14.7667 | 10.6667 | 14.5 | 14.8433 | 11.9183 | 11.3333 | 12.1058 | 15.4417 | 15.7672 | 17.9167 | 15.4333 | 14.2767 | 15.2167
| Median |  | 16.45 | 14.8 | 13.8668 | 13.1785 | 12.25 | 14.5 | 14.95 | 14.93 | 18 | 14 | 14.175 | 15.475 | 14.525 | 14.725 | 11 | 14.55 | 14.845 | 12.55 | 11 | 12.0242 | 15.35 | 15.8044 | 17.9 | 15.45 | 14.345 | 15.35
| Std Dev. |  | 0.52915 | 0.061968 | 0.21534 | 0.16609 | 0.53166 | 0.96264 | 0.28577 | 0.30725 | 0.5164 | 0.40825 | 0.16733 | 0.70905 | 0.22361 | 0.27325 | 0.5164 | 0.85557 | 0.27398 | 2.1003 | 0.5164 | 0.68163 | 0.16558 | 0.60355 | 0.31885 | 0.50067 | 0.30111 | 0.86814
| Rel.Std.Dev. |  | 3.246% | 0.4184% | 1.554% | 1.258% | 4.299% | 6.608% | 1.907% | 2.055% | 2.923% | 2.882% | 1.178% | 4.627% | 1.537% | 1.85% | 4.841% | 5.9% | 1.846% | 17.62% | 4.556% | 5.631% | 1.072% | 3.828% | 1.78% | 3.244% | 2.109% | 5.705%
| PDM3 |  | 9.892% | -0.1538% | -6.565% | -11% | -16.63% | -1.794% | 1.015% | 0.7901% | 19.11% | -4.491% | -4.266% | 3.318% | -1.907% | -0.4459% | -28.09% | -2.244% | 0.07095% | -19.65% | -23.59% | -18.38% | 4.105% | 6.299% | 20.79% | 4.049% | -3.749% | 2.588%
| Z-Score (Absolute) |  | 1.81 | 0.03 | 1.11 | 1.9 | 2.9 | 0.27 | 0.23 | 0.19 | NA | NA | 0.71 | 0.64 | 0.29 | 0.03 | NA | 0.35 | 0.07 | 3.44 | NA | 3.22 | 0.78 | 1.17 | 3.75 | 0.77 | 0.61 | 0.51
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 240b (Certified Value 2.89 ppm) | [30 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.04 | 2.52 | 3.3524 | 2.7 | 2.4 | 2.91 | 3.1 | 0.51867 | [21 empty] | 1
| 1 | 2 | 2.98 | 2.68 | 3.3387 | 2.8 | 2.7 | 2.94 | 2.8 | 0.50421 | [21 empty] | 33
| 1 | 3 | 3.03 | 2.7 | 3.3929 | 2.7 | 2.6 | 3.06 | 2.8 | 0.53996 | [21 empty] | 16
| 1 | 4 | 3.07 | 2.52 | 3.3339 | 2.7 | 2.6 | 2.96 | 2.8 | 0.52788 | [21 empty] | 2.8855
| 1 | 5 | 2.87 | 2.55 | 3.2666 | 2.9 | 2.7 | 3.04 | 2.9 | 0.52698 | [21 empty] | 33
| 1 | 6 | 3.09 | 2.71 | 3.3181 | 2.7 | 3 | 3.1 | 2.8 | 0.5413
| Mean |  | 3.0133 | 2.6133 | 3.3338 | 2.75 | 2.6667 | 3.0017 | 2.8667 | 0.5265
| Median |  | 3.035 | 2.615 | 3.3363 | 2.7 | 2.65 | 3 | 2.8 | 0.52743
| Std Dev. |  | 0.079666 | 0.092448 | 0.041526 | 0.083666 | 0.19664 | 0.075476 | 0.12111 | 0.01386
| Rel.Std.Dev. |  | 2.644% | 3.538% | 1.246% | 3.042% | 7.374% | 2.514% | 4.225% | 2.632%
| PDM3 |  | 4.429% | -9.434% | 15.53% | -4.697% | -7.585% | 4.024% | -0.6541% | -81.75%
| Z-Score (Absolute) |  | 0.71 | 0.68 | 1.82 | 0.2 | 0.49 | 0.67 | 0.2 | 7.92
| Analytical results for Ge in OREAS 240b (Certified Value 0.095 ppm) | [30 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | < 0.1 | 1.7 | 0.1471 | < 0.2 | 0.2 | 1.5 | 1.3 | 2.8 | 0.09 | 0.1 | 0.07 | 0.08 | < 10 | 0.6 | < 5 | 0.06 | 0.0903 | 0.07 | 0.2 | 0.06 | 0.15 | [8 empty] | 1
| 1 | 2 | < 0.1 | 1.8 | 0.15996 | < 0.2 | 0.2 | 1.7 | 1.2 | 2.7 | 0.09 | 0.1 | 0.08 | 0.06 | < 10 | 0.6 | < 5 | 0.06 | 0.1074 | 0.08 | 0.2 | < 0.05 | 0.12 | [8 empty] | 34
| 1 | 3 | < 0.1 | 1.7 | 0.1503 | < 0.2 | 0.3 | 2 | 1.3 | 2.8 | 0.08 | 0.08 | 0.08 | 0.08 | < 10 | 0.6 | < 5 | 0.07 | 0.0991 | 0.11 | 0.2 | 0.06 | 0.12 | [8 empty] | 16
| 1 | 4 | < 0.1 | 1.7 | 0.15398 | < 0.2 | 0.3 | 2 | 1.1 | 2.9 | 0.08 | 0.09 | 0.07 | 0.07 | < 10 | 0.7 | < 5 | 0.07 | 0.14 | 0.1 | 0.1 | < 0.05 | 0.13 | [8 empty] | 0.0953
| 1 | 5 | < 0.1 | 1.6 | 0.15851 | < 0.2 | 0.3 | 1.7 | 1.2 | 3 | 0.09 | 0.1 | 0.09 | 0.06 | < 10 | 0.7 | < 5 | 0.06 | 0.12666 | 0.1 | 0.1 | < 0.05 | 0.11 | [8 empty] | 34
| 1 | 6 | < 0.1 | 1.7 | 0.159 | < 0.2 | 0.3 | 1.8 | 1.3 | 2.9 | 0.09 | 0.12 | 0.07 | 0.07 | < 10 | 0.7 | < 5 | 0.07 | 0.1757 | 0.13 | 0.2 | < 0.05 | 0.12
| Mean |  |  | 1.7 | 0.15481 |  | 0.26667 | 1.7833 | 1.2333 | 2.85 | 0.086667 | 0.098333 | 0.076667 | 0.07 |  | 0.65 |  | 0.065 | 0.12319 | 0.098333 | 0.16667 | 0.06 | 0.125
| Median |  |  | 1.7 | 0.15625 |  | 0.3 | 1.75 | 1.25 | 2.85 | 0.09 | 0.1 | 0.075 | 0.07 |  | 0.65 |  | 0.065 | 0.11703 | 0.1 | 0.2 | 0.06 | 0.12
| Std Dev. |  |  | 0.063246 | 0.0052596 |  | 0.05164 | 0.19408 | 0.08165 | 0.10488 | 0.005164 | 0.013292 | 0.008165 | 0.0089443 |  | 0.054772 |  | 0.0054772 | 0.031507 | 0.02137 | 0.05164 | 0 | 0.013784
| Rel.Std.Dev. |  |  | 3.72% | 3.397% |  | 19.36% | 10.88% | 6.62% | 3.68% | 5.958% | 13.52% | 10.65% | 12.78% |  | 8.427% |  | 8.427% | 25.57% | 21.73% | 30.98% | 0% | 11.03%
| PDM3 |  |  | 1684% | 62.44% |  | 179.8% | 1771% | 1194% | 2891% | -9.059% | 3.183% | -19.55% | -26.55% |  | 582.1% |  | -31.79% | 29.27% | 3.183% | 74.89% | -37.04% | 31.16%
| Z-Score (Absolute) |  | 1.19 | NA | 1.38 | 0.04 | NA | NA | NA | NA | 0.29 | 0 | 0.53 | 0.69 | 120.19 | NA | 58.89 | 0.82 | 0.61 | 0 | NA | 1.51 | 0.65
Results from laboratories 1.02, 1.07, 1.10, 1.11, 1.13, 1.22 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 240b (Certified Value 1.66 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.89 | 1.49 | 0.72434 | 1.4 | 1.5 | 1.9 | 1.76 | 1.9 | 1.6 | 1.7 | 1.7 | 1.5 | 1.6 | 1.4 | 1.71 | 2 | 1.9151 | 1.7 | 1.5106 | 1.8 | 2.32 | 1.8 | 1.69 | [6 empty] | 1
| 1 | 2 | 1.79 | 1.6 | 0.71716 | 0.7 | 1.7 | 1.7 | 1.72 | 1.8 | 1.6 | 1.7 | 1.7 | 1.7 | 1.7 | 1.5 | 1.69 | 1.5 | 1.9354 | 1.7 | 1.5553 | 1.9 | 2.24 | 1.8 | 1.62 | [6 empty] | 18
| 1 | 3 | 1.88 | 1.56 | 0.71661 | 1.1 | 1.6 | 1.9 | 1.69 | 1.7 | 1.5 | 1.7 | 1.7 | 1.5 | 1.7 | 1.5 | 1.7 | 1.5 | 2.0526 | 1.7 | 1.5219 | 1.8 | 2.36 | 1.9 | 1.53 | [6 empty] | 16
| 1 | 4 | 1.78 | 1.52 | 0.72722 | 1.6 | 1.7 | 1.6 | 1.58 | 1.7 | 1.5 | 1.7 | 1.6 | 1.8 | 1.5 | 1.4 | 1.7 | 1.5 | 1.9805 | 1.6 | 1.5264 | 1.9 | 2.27 | 1.9 | 1.67 | [6 empty] | 1.6641
| 1 | 5 | 1.61 | 1.47 | 0.69205 | 1.9 | 1.7 | 2 | 1.64 | 1.8 | 1.6 | 1.6 | 1.5 | 1.5 | 1.5 | 1.4 | 1.71 | 1.5 | 1.9972 | 1.6 | 1.4915 | 1.8 | 2.41 | 1.7 | 1.51 | [6 empty] | 35
| 1 | 6 | 1.68 | 1.74 | 0.66357 | 1.7 | 1.9 | 1.6 | 1.63 | 1.7 | 1.5 | 1.4 | 1.6 | 1.5 | 1.6 | 1.4 | 1.73 | 1.5 | 2.0393 | 1.7 | 1.4584 | 1.8 | 2.28 | 1.8 | 1.58
| Mean |  | 1.7717 | 1.5633 | 0.70683 | 1.4 | 1.6833 | 1.7833 | 1.67 | 1.7667 | 1.55 | 1.6333 | 1.6333 | 1.5833 | 1.6 | 1.4333 | 1.7067 | 1.5833 | 1.9867 | 1.6667 | 1.5107 | 1.8333 | 2.3133 | 1.8167 | 1.6
| Median |  | 1.785 | 1.54 | 0.71689 | 1.5 | 1.7 | 1.8 | 1.665 | 1.75 | 1.55 | 1.7 | 1.65 | 1.5 | 1.6 | 1.4 | 1.705 | 1.5 | 1.9888 | 1.7 | 1.5162 | 1.8 | 2.3 | 1.8 | 1.6
| Std Dev. |  | 0.11017 | 0.098522 | 0.024554 | 0.43818 | 0.13292 | 0.17224 | 0.065727 | 0.08165 | 0.054772 | 0.12111 | 0.08165 | 0.13292 | 0.089443 | 0.05164 | 0.013663 | 0.20412 | 0.05481 | 0.05164 | 0.033061 | 0.05164 | 0.06314 | 0.075277 | 0.073212
| Rel.Std.Dev. |  | 6.218% | 6.302% | 3.474% | 31.3% | 7.896% | 9.658% | 3.936% | 4.622% | 3.534% | 7.415% | 4.999% | 8.395% | 5.59% | 3.603% | 0.8005% | 12.89% | 2.759% | 3.098% | 2.188% | 2.817% | 2.729% | 4.144% | 4.576%
| PDM3 |  | 6.465% | -6.054% | -57.52% | -15.87% | 1.157% | 7.166% | 0.3557% | 6.165% | -6.855% | -1.848% | -1.848% | -4.852% | -3.851% | -13.87% | 2.559% | -4.852% | 19.39% | 0.1554% | -9.218% | 10.17% | 39.02% | 9.169% | -3.851%
| Z-Score (Absolute) |  | 1.12 | 0.57 | 7.5 | 1.89 | 0.4 | 1.21 | 0.3 | 1.08 | 0.67 | 0 | 0 | 0.4 | 0.27 | 1.62 | 0.59 | 0.4 | 2.86 | 0.27 | 0.99 | 1.62 | 5.5 | 1.48 | 0.27
| Analytical results for Hg in OREAS 240b (Indicative Value 0.098 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.05 | 1.25 | 1.30 | [26 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 3
3
| 1 | 1 | < 1 | < 0.1 | 0.1203 | [26 empty] | 1
| 1 | 2 | < 1 | < 0.1 | 0.1243 | [26 empty] | 5
| 1 | 3 | < 1 | < 0.1 | 0.0984 | [26 empty] | 16
| 1 | 4 | < 1 | < 0.1 | 0.0987 | [26 empty] | 0.098133
| 1 | 5 | < 1 | < 0.1 | 0.0869 | [26 empty] | 9
| 1 | 6 | < 1 | < 0.1 | 0.0602
| Mean |  |  |  | 0.098133
| Median |  |  |  | 0.09855
| Std Dev. |  |  |  | 0.023416
| Rel.Std.Dev. |  |  |  | 23.86%
| PDM3 |  |  |  | -5.551e-14%
| Z-Score (Absolute) |  | 5.63 | 0.67 | 0
| Analytical results for Ho in OREAS 240b (Certified Value 0.67 ppm) | [30 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.72 | 0.61 | 0.67089 | 0.7 | 0.6 | 0.68 | 0.65 | 0.71963 | [21 empty] | 1
| 1 | 2 | 0.71 | 0.66 | 0.65077 | 0.7 | 0.66 | 0.7 | 0.7 | 0.70342 | [21 empty] | 19
| 1 | 3 | 0.72 | 0.66 | 0.63707 | 0.7 | 0.64 | 0.69 | 0.65 | 0.78209 | [21 empty] | 16
| 1 | 4 | 0.71 | 0.62 | 0.65733 | 0.7 | 0.64 | 0.7 | 0.65 | 0.73884 | [21 empty] | 0.67009
| 1 | 5 | 0.72 | 0.62 | 0.64159 | 0.7 | 0.66 | 0.71 | 0.7 | 0.73514 | [21 empty] | 36
| 1 | 6 | 0.75 | 0.66 | 0.64946 | 0.7 | 0.74 | 0.72 | 0.6 | 0.76276
| Mean |  | 0.72167 | 0.63833 | 0.65118 | 0.7 | 0.65667 | 0.7 | 0.65833 | 0.74031
| Median |  | 0.72 | 0.64 | 0.65012 | 0.7 | 0.65 | 0.7 | 0.65 | 0.73699
| Std Dev. |  | 0.01472 | 0.024014 | 0.012003 | 1.2162e-16 | 0.046332 | 0.014142 | 0.037639 | 0.02852
| Rel.Std.Dev. |  | 2.04% | 3.762% | 1.843% | 1.737e-14% | 7.056% | 2.02% | 5.717% | 3.852%
| PDM3 |  | 7.698% | -4.739% | -2.821% | 4.464% | -2.003% | 4.464% | -1.754% | 10.48%
| Z-Score (Absolute) |  | 2.14 | 0.67 | 0.24 | NA | 0.06 | 1.4 | 0 | 2.76
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 240b (Certified Value 0.075 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.08 | 0.08 | 0.06161 | < 0.1 | 0.08 | 0.07 | 0.08 | 0.078 | 0.073 | 0.074 | 0.077 | < 10 | 0.06 | 0.07 | 0.074 | 30.61 | < 0.1 | 0.06509 | 0.058 | 0.0834 | 0.072 | 0.09 | 0.076 | 0.081 | [5 empty] | 1
| 1 | 2 | 0.09 | 0.08 | 0.06161 | < 0.1 | 0.08 | 0.07 | 0.07 | 0.079 | 0.076 | 0.076 | 0.072 | < 10 | 0.07 | 0.07 | 0.074 | 13.02 | < 0.1 | 0.0617 | 0.046 | 0.0861 | 0.072 | 0.08 | 0.069 | 0.079 | [5 empty] | 20
| 1 | 3 | 0.08 | 0.08 | 0.057572 | < 0.1 | 0.08 | 0.09 | 0.08 | 0.072 | 0.076 | 0.075 | 0.076 | < 10 | 0.06 | 0.07 | 0.075 | 31.85 | < 0.1 | 0.07279 | 0.058 | 0.0859 | 0.072 | 0.08 | 0.073 | 0.073 | [5 empty] | 16
| 1 | 4 | 0.09 | 0.07 | 0.05453 | < 0.1 | 0.08 | 0.07 | 0.08 | 0.072 | 0.075 | 0.077 | 0.083 | < 10 | 0.07 | 0.07 | 0.077 | 28.6 | < 0.1 | 0.06596 | 0.047 | 0.0781 | 0.076 | 0.08 | 0.078 | 0.081 | [5 empty] | 0.075469
| 1 | 5 | 0.08 | 0.06 | 0.060242 | < 0.1 | 0.07 | 0.08 | 0.08 | 0.076 | 0.075 | 0.074 | 0.072 | < 10 | 0.07 | 0.07 | 0.076 | 33.23 | < 0.1 | 0.06741 | 0.056 | 0.0748 | 0.086 | 0.07 | 0.073 | 0.075 | [5 empty] | 37
| 1 | 6 | 0.09 | 0.08 | 0.06136 | < 0.1 | 0.08 | 0.07 | 0.08 | 0.079 | 0.084 | 0.079 | 0.075 | < 10 | 0.07 | 0.07 | 0.076 | 26.16 | < 0.1 | 0.0693 | 0.055 | 0.0739 | 0.074 | 0.08 | 0.072 | 0.079
| Mean |  | 0.085 | 0.075 | 0.059487 |  | 0.078333 | 0.075 | 0.078333 | 0.076 | 0.0765 | 0.075833 | 0.075833 |  | 0.066667 | 0.07 | 0.075333 | 27.245 |  | 0.067042 | 0.053333 | 0.080367 | 0.075333 | 0.08 | 0.0735 | 0.078
| Median |  | 0.085 | 0.08 | 0.060801 |  | 0.08 | 0.07 | 0.08 | 0.077 | 0.0755 | 0.0755 | 0.0755 |  | 0.07 | 0.07 | 0.0755 | 29.605 |  | 0.066685 | 0.0555 | 0.08075 | 0.073 | 0.08 | 0.073 | 0.079
| Std Dev. |  | 0.0054772 | 0.0083666 | 0.0028753 |  | 0.0040825 | 0.0083666 | 0.0040825 | 0.0032863 | 0.0038341 | 0.0019408 | 0.0040702 |  | 0.005164 | 0 | 0.0012111 | 7.3988 |  | 0.0037892 | 0.0054283 | 0.0054888 | 0.005465 | 0.0063246 | 0.0031464 | 0.0032863
| Rel.Std.Dev. |  | 6.444% | 11.16% | 4.833% |  | 5.212% | 11.16% | 5.212% | 4.324% | 5.012% | 2.559% | 5.367% |  | 7.746% | 0% | 1.608% | 27.16% |  | 5.652% | 10.18% | 6.83% | 7.254% | 7.906% | 4.281% | 4.213%
| PDM3 |  | 12.63% | -0.6214% | -21.18% |  | 3.795% | -0.6214% | 3.795% | 0.7036% | 1.366% | 0.4828% | 0.4828% |  | -11.66% | -7.247% | -0.1797% | 36001% |  | -11.17% | -29.33% | 6.49% | -0.1797% | 6.004% | -2.609% | 3.354%
| Z-Score (Absolute) |  | 1.77 | 0.11 | 3.03 | 4.81 | 0.52 | 0.11 | 0.52 | 0.08 | 0.17 | 0.05 | 0.05 | 926.67 | 1.68 | 1.05 | 0.05 | 5112.72 | 4.81 | 1.61 | 4.19 | 0.9 | 0.05 | 0.83 | 0.39 | 0.45
| Analytical results for K in OREAS 240b (Certified Value 0.73 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.75 | 0.7089 | 0.7253 | 0.72344 | 0.72 | 0.7704 | 0.717 | 0.682 | 0.73 | 0.75 | 0.72 | 0.72 | 0.69 | 0.68 | 0.77 | 0.721 | 0.69 | 0.76 | 1.36 | 0.71664 | 0.7 | 0.71 | 0.77881 | 0.72 | 0.8 | 0.76 | 0.72 |  |  | 1
| 1 | 2 | 0.76 | 0.7148 | 0.72446 | 0.72714 | 0.68 | 0.7729 | 0.72 | 0.6828 | 0.73 | 0.75 | 0.72 | 0.75 | 0.71 | 0.68 | 0.74 | 0.74 | 0.67 | 0.74 | 1.646 | 0.71487 | 0.705 | 0.72 | 0.7938 | 0.74 | 0.77 | 0.75 | 0.75 |  |  | #N/A
| 1 | 3 | 0.77 | 0.7207 | 0.72212 | 0.72835 | 0.72 | 0.7712 | 0.722 | 0.723 | 0.71 | 0.74 | 0.71 | 0.75 | 0.7 | 0.69 | 0.76 | 0.733 | 0.68 | 0.72 | 1.045 | 0.73317 | 0.695 | 0.73 | 0.79951 | 0.73 | 0.79 | 0.76 | 0.75 |  |  | 16
| 1 | 4 | 0.76 | 0.7218 | 0.72141 | 0.71233 | 0.72 | 0.782 | 0.724 | 0.702 | 0.72 | 0.75 | 0.7 | 0.74 | 0.7 | 0.69 | 0.75 | 0.733 | 0.66 | 0.73 | 1.322 | 0.73742 | 0.695 | 0.73 | 0.79491 | 0.73 | 0.79 | 0.76 | 0.71 |  |  | 0.72952
| 1 | 5 | 0.74 | 0.6948 | 0.72255 | 0.70929 | 0.74 | 0.7729 | 0.713 | 0.6942 | 0.75 | 0.77 | 0.73 | 0.75 | 0.68 | 0.7 | 0.74 | 0.724 | 0.65 | 0.77 | 1.365 | 0.74664 | 0.695 | 0.72 | 0.77421 | 0.74 | 0.75 | 0.75 | 0.76 |  |  | 38
| 1 | 6 | 0.76 | 0.7246 | 0.72925 | 0.71387 | 0.69 | 0.7795 | 0.709 | 0.6786 | 0.73 | 0.77 | 0.73 | 0.71 | 0.68 | 0.7 | 0.73 | 0.734 | 0.66 | 0.75 | 1.189 | 0.7392 | 0.7 | 0.73 | 0.76971 | 0.73 | 0.77 | 0.75 | 0.74
| Mean |  | 0.75667 | 0.71427 | 0.72418 | 0.71907 | 0.71167 | 0.77482 | 0.7175 | 0.69377 | 0.72833 | 0.755 | 0.71833 | 0.73667 | 0.69333 | 0.69 | 0.74833 | 0.73083 | 0.66833 | 0.745 | 1.3212 | 0.73132 | 0.69833 | 0.72333 | 0.78516 | 0.73167 | 0.77833 | 0.755 | 0.73833
| Median |  | 0.76 | 0.71775 | 0.7235 | 0.71866 | 0.72 | 0.7729 | 0.7185 | 0.6885 | 0.73 | 0.75 | 0.72 | 0.745 | 0.695 | 0.69 | 0.745 | 0.733 | 0.665 | 0.745 | 1.341 | 0.7353 | 0.6975 | 0.725 | 0.78631 | 0.73 | 0.78 | 0.755 | 0.745
| Std Dev. |  | 0.010328 | 0.011072 | 0.0028821 | 0.0082275 | 0.022286 | 0.0047638 | 0.0056833 | 0.016801 | 0.013292 | 0.012247 | 0.01169 | 0.017512 | 0.012111 | 0.0089443 | 0.01472 | 0.0070261 | 0.01472 | 0.018708 | 0.20134 | 0.012832 | 0.0040825 | 0.008165 | 0.012447 | 0.0075277 | 0.018348 | 0.0054772 | 0.019408
| Rel.Std.Dev. |  | 1.365% | 1.55% | 0.398% | 1.144% | 3.132% | 0.6148% | 0.7921% | 2.422% | 1.825% | 1.622% | 1.627% | 2.377% | 1.747% | 1.296% | 1.967% | 0.9614% | 2.202% | 2.511% | 15.24% | 1.755% | 0.5846% | 1.129% | 1.585% | 1.029% | 2.357% | 0.7255% | 2.629%
| PDM3 |  | 3.721% | -2.091% | -0.7321% | -1.433% | -2.448% | 6.209% | -1.648% | -4.901% | -0.1629% | 3.492% | -1.534% | 0.9794% | -4.961% | -5.418% | 2.579% | 0.1797% | -8.388% | 2.122% | 81.1% | 0.247% | -4.275% | -0.8483% | 7.626% | 0.294% | 6.691% | 3.492% | 1.208%
| Z-Score (Absolute) |  | 1 | 0.64 | 0.26 | 0.45 | 0.74 | 1.69 | 0.51 | 1.43 | 0.1 | 0.93 | 0.48 | 0.22 | 1.44 | 1.57 | 0.67 | 0 | 2.41 | 0.55 | 22.75 | 0.02 | 1.25 | 0.29 | 2.09 | 0.03 | 1.83 | 0.93 | 0.29
| Analytical results for La in OREAS 240b (Certified Value 6.6 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 6.8 | 6.47 | 5.96 | 6.73 | 6.5291 | 6.2 | 7 | 6.24 | 6 | 9 | 6.5 | 7.1 | 6.5 | 6.4 | 7 | 6.2 | 6.2 | 7.5 | 6.4285 | 6.9 | 6.4672 | 8.1 | 6.8 | 7.3 | 7.9 |  |  |  |  | 1
| 1 | 2 | 6.8 | 6.56 | 6.26 | 6.83 | 6.7011 | 6.3 | 7 | 6.36 | 6 | 10 | 6.4 | 7.4 | 7 | 6.4 | 7 | 6.2 | 6.3 | 7.5 | 6.2898 | 6.7 | 6.5708 | 7.9 | 6.7 | 7.1 | 7.6 |  |  |  |  | 21
| 1 | 3 | 7 | 6.64 | 6.25 | 6.77 | 6.6965 | 6.3 | 7 | 6.55 | 6 | 9 | 6.3 | 7.1 | 6.6 | 6.6 | 7 | 6.2 | 6.3 | 7 | 6.7123 | 7.1 | 6.5733 | 8 | 6.9 | 7.3 | 7.1 |  |  |  |  | 16
| 1 | 4 | 7 | 6.45 | 5.89 | 6.78 | 6.6206 | 6.4 | 7 | 6.39 | 6 | 9 | 6.1 | 6.8 | 6.5 | 6.7 | 7 | 5.9 | 6.5 | 7 | 6.6075 | 6.9 | 6.4073 | 7.9 | 6.6 | 7.5 | 8.1 |  |  |  |  | 6.6049
| 1 | 5 | 6.1 | 6.32 | 5.92 | 6.64 | 6.7127 | 6.9 | 7 | 6.34 | 7 | 10 | 6.2 | 7.2 | 6.5 | 6.4 | 7 | 6.1 | 6.4 | 7.5 | 6.4719 | 6.7 | 6.228 | 8 | 6.3 | 7.3 | 6.9 |  |  |  |  | 39
| 1 | 6 | 7 | 6.71 | 6.26 | 6.69 | 6.6436 | 6.3 | 7 | 6.26 | 6 | 10 | 6.3 | 7.1 | 6.5 | 6.2 | 7 | 6.1 | 6.2 | 7.5 | 6.8826 | 6.8 | 6.0071 | 8.1 | 6.6 | 7.4 | 7.7
| Mean |  | 6.7833 | 6.525 | 6.09 | 6.74 | 6.6506 | 6.4 | 7 | 6.3567 | 6.1667 | 9.5 | 6.3 | 7.1167 | 6.6 | 6.45 | 7 | 6.1167 | 6.3167 | 7.3333 | 6.5654 | 6.85 | 6.3756 | 8 | 6.65 | 7.3167 | 7.55
| Median |  | 6.9 | 6.515 | 6.105 | 6.75 | 6.67 | 6.3 | 7 | 6.35 | 6 | 9.5 | 6.3 | 7.1 | 6.5 | 6.4 | 7 | 6.15 | 6.3 | 7.5 | 6.5397 | 6.85 | 6.4373 | 8 | 6.65 | 7.3 | 7.65
| Std Dev. |  | 0.34881 | 0.14096 | 0.18396 | 0.068118 | 0.069604 | 0.25298 | 0 | 0.11112 | 0.40825 | 0.54772 | 0.14142 | 0.19408 | 0.2 | 0.17607 | 0 | 0.1169 | 0.1169 | 0.2582 | 0.21326 | 0.15166 | 0.22098 | 0.089443 | 0.20736 | 0.13292 | 0.46368
| Rel.Std.Dev. |  | 5.142% | 2.16% | 3.021% | 1.011% | 1.047% | 3.953% | 0% | 1.748% | 6.62% | 5.766% | 2.245% | 2.727% | 3.03% | 2.73% | 0% | 1.911% | 1.851% | 3.521% | 3.248% | 2.214% | 3.466% | 1.118% | 3.118% | 1.817% | 6.141%
| PDM3 |  | 2.701% | -1.21% | -7.796% | 2.045% | 0.6918% | -3.102% | 5.982% | -3.758% | -6.635% | 43.83% | -4.616% | 7.748% | -0.07437% | -2.345% | 5.982% | -7.392% | -4.364% | 11.03% | -0.5975% | 3.711% | -3.472% | 21.12% | 0.6826% | 10.78% | 14.31%
| Z-Score (Absolute) |  | 0.51 | 0.21 | 1.41 | 0.39 | 0.14 | 0.55 | NA | 0.67 | NA | NA | 0.83 | 1.43 | 0 | 0.42 | NA | 1.34 | 0.79 | 2.03 | 0.1 | 0.69 | 0.62 | 3.88 | 0.14 | 1.99 | 2.63
Results from laboratories 1.08, 1.12, 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 240b (Certified Value 13.5 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | < 1 | 13.1 | 15.89 | 7.1278 | 14.1042 | 13 | 8 | 14.1 | 12.9 | 13 | 15.9 | 13.6 | 13.7 | 13.8 | 13.5 | < 5 | 13.3 | 11.9 | 13 | 14 | 10.7965 | 13.3 | 11.6142 | 13.9 | 16 | 11.8 | 13.5 |  |  | 1
| 1 | 2 | < 1 | 13.2 | 14.22 | 7.0294 | 14.1725 | 13 | 8 | 14.3 | 12.7 | 13 | 16.2 | 13.9 | 14.3 | 14.4 | 13.4 | < 5 | 13.4 | 11.7 | 13 | 14 | 10.2878 | 13.5 | 11.7624 | 13.8 | 16 | 11.2 | 13.8 |  |  | #N/A
| 1 | 3 | < 1 | 13.3 | 14.35 | 7.1461 | 14.0498 | 13 | 8 | 14.2 | 13.3 | 12 | 13.8 | 13.8 | 14.2 | 13.8 | 13.4 | < 5 | 13.3 | 11.9 | 12 | 13 | 11.6604 | 13.6 | 11.7468 | 13.7 | 16 | 11.5 | 12.2 |  |  | 16
| 1 | 4 | < 1 | 13.2 | 13.4 | 7.0995 | 13.9307 | 13 | 8 | 14.2 | 12.7 | 13 | 14 | 13.8 | 14 | 13.8 | 13.7 | < 5 | 13.4 | 11.6 | 13 | 14 | 10.9715 | 13.7 | 11.6922 | 13.5 | 17 | 11.4 | 13.7 |  |  | 13.4658
| 1 | 5 | < 1 | 12.9 | 13.62 | 6.9758 | 13.7817 | 13 | 8 | 14.3 | 12.2 | 13 | 14.7 | 14 | 14 | 13.4 | 13.8 | < 5 | 13.3 | 11.5 | 14 | 14 | 10.9718 | 13.3 | 11.5908 | 13.7 | 15 | 11.1 | 12.9 |  |  | 40
| 1 | 6 | < 1 | 13.5 | 13.38 | 7.1445 | 13.905 | 13 | 8 | 14.3 | 12.6 | 13 | 14.4 | 14.2 | 13.7 | 13.6 | 13.6 | < 5 | 13.3 | 11.5 | 13 | 14 | 11.433 | 13.5 | 11.505 | 13.8 | 16 | 11.2 | 13.1
| Mean |  |  | 13.2 | 14.1433 | 7.0872 | 13.9906 | 13 | 8 | 14.2333 | 12.7333 | 12.8333 | 14.8333 | 13.8833 | 13.9833 | 13.8 | 13.5667 |  | 13.3333 | 11.6833 | 13 | 13.8333 | 11.0202 | 13.4833 | 11.6519 | 13.7333 | 16 | 11.3667 | 13.2
| Median |  |  | 13.2 | 13.92 | 7.1136 | 13.9902 | 13 | 8 | 14.25 | 12.7 | 13 | 14.55 | 13.85 | 14 | 13.8 | 13.55 |  | 13.3 | 11.65 | 13 | 14 | 10.9716 | 13.5 | 11.6532 | 13.75 | 16 | 11.3 | 13.3
| Std Dev. |  |  | 0.2 | 0.94958 | 0.069747 | 0.14427 | 0 | 0 | 0.08165 | 0.36148 | 0.40825 | 0.99733 | 0.20412 | 0.24833 | 0.33466 | 0.1633 |  | 0.05164 | 0.18348 | 0.63246 | 0.40825 | 0.48404 | 0.16021 | 0.099574 | 0.13663 | 0.63246 | 0.2582 | 0.6
| Rel.Std.Dev. |  |  | 1.515% | 6.714% | 0.9841% | 1.031% | 0% | 0% | 0.5737% | 2.839% | 3.181% | 6.724% | 1.47% | 1.776% | 2.425% | 1.204% |  | 0.3873% | 1.57% | 4.865% | 2.951% | 4.392% | 1.188% | 0.8546% | 0.9948% | 3.953% | 2.272% | 4.545%
| PDM3 |  |  | -1.974% | 5.032% | -47.37% | 3.898% | -3.459% | -40.59% | 5.7% | -5.439% | -4.697% | 10.16% | 3.101% | 3.843% | 2.482% | 0.7492% |  | -0.9836% | -13.24% | -3.459% | 2.73% | -18.16% | 0.1303% | -13.47% | 1.987% | 18.82% | -15.59% | -1.974%
| Z-Score (Absolute) |  | 13.17 | 0.14 | 0.83 | 6.41 | 0.67 | NA | NA | 0.92 | 0.62 | NA | 1.54 | 0.56 | 0.67 | 0.48 | 0.24 | 11.11 | 0 | 1.69 | NA | NA | 2.37 | 0.15 | 1.72 | 0.41 | NA | 2.02 | 0.14
Results from laboratories 1.07, 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 240b (Certified Value 0.29 ppm) | [30 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | 1.32 | [20 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [20 empty] | 2
3
| 1 | 1 | 0.32 | 0.28 | 0.25 | 4.007 | 0.3 | 0.3 | 0.4 | 0.29305 | 0.29 | [20 empty] | 1
| 1 | 2 | 0.32 | 0.29 | 0.26 | 4.0329 | 0.3 | 0.3 | 0.4 | 0.30761 | 0.28 | [20 empty] | 22
| 1 | 3 | 0.33 | 0.32 | 0.26 | 4.0762 | 0.3 | 0.31 | 0.4 | 0.32643 | 0.29 | [20 empty] | 16
| 1 | 4 | 0.33 | 0.3 | 0.25 | 4.0633 | 0.3 | 0.29 | 0.5 | 0.31755 | 0.28 | [20 empty] | 0.2946
| 1 | 5 | 0.31 | 0.28 | 0.25 | 4.0111 | 0.3 | 0.31 | 0.4 | 0.30864 | 0.26 | [20 empty] | 41
| 1 | 6 | 0.33 | 0.3 | 0.26 | 4.0992 | 0.3 | 0.31 | 0.4 | 0.32233 | 0.27
| Mean |  | 0.32333 | 0.295 | 0.255 | 4.0483 | 0.3 | 0.30333 | 0.41667 | 0.3126 | 0.27833
| Median |  | 0.325 | 0.295 | 0.255 | 4.0481 | 0.3 | 0.305 | 0.4 | 0.3131 | 0.28
| Std Dev. |  | 0.008165 | 0.015166 | 0.0054772 | 0.037191 | 0 | 0.008165 | 0.040825 | 0.01211 | 0.01169
| Rel.Std.Dev. |  | 2.525% | 5.141% | 2.148% | 0.9187% | 0% | 2.692% | 9.798% | 3.874% | 4.2%
| PDM3 |  | 9.753% | 0.1357% | -13.44% | 1274% | 1.833% | 2.964% | 41.43% | 6.11% | -5.522%
| Z-Score (Absolute) |  | 0.67 | 0.28 | 1.63 | 126.26 | NA | 0 | NA | 0.31 | 0.84
Results from laboratories 1.07 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 240b (Certified Value 3.44 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.34 | 3.293 | 3.621 | 3.4334 | 3.19 | 3.6666 | 3.53 | 3.2904 | 3.27 | 3.1 | 3.45 | 3.42 | 3.42 | 3.38 | 3.61 | 3.48 | 3.46 | 3.392 | 1.18 | 2.9068 | 3.5885 | 3.43 | 3.532 | 3.48 | 2.82 | 3.49 | 3.9 |  |  | 1
| 1 | 2 | 3.39 | 3.2869 | 3.648 | 3.4463 | 3.14 | 3.6847 | 3.54 | 3.3052 | 3.28 | 3.1 | 3.48 | 3.58 | 3.5 | 3.36 | 3.5 | 3.548 | 3.45 | 3.33 | 0.57 | 2.9022 | 3.585 | 3.53 | 3.5731 | 3.56 | 2.72 | 3.46 | 3.6 |  |  | #N/A
| 1 | 3 | 3.4 | 3.326 | 3.6653 | 3.4337 | 3.22 | 3.6425 | 3.54 | 3.4751 | 3.27 | 3.08 | 3.43 | 3.56 | 3.4 | 3.35 | 3.59 | 3.549 | 3.47 | 3.277 | 1.75 | 2.9714 | 3.5635 | 3.51 | 3.562 | 3.47 | 2.81 | 3.49 | 3 |  |  | 16
| 1 | 4 | 3.36 | 3.2815 | 3.591 | 3.3901 | 3.2 | 3.733 | 3.55 | 3.3362 | 3.25 | 3.04 | 3.43 | 3.53 | 3.42 | 3.43 | 3.55 | 3.543 | 3.42 | 3.302 | 0.63 | 3.0296 | 3.565 | 3.51 | 3.543 | 3.56 | 2.81 | 3.47 | 3.7 |  |  | 3.4378
| 1 | 5 | 3.28 | 3.3163 | 3.6602 | 3.3414 | 3.45 | 3.6666 | 3.49 | 3.2575 | 3.29 | 3.15 | 3.54 | 3.57 | 3.3 | 3.42 | 3.51 | 3.488 | 3.37 | 3.443 | 1.44 | 3.0335 | 3.5655 | 3.51 | 3.4886 | 3.56 | 2.69 | 3.47 | 2.9 |  |  | 42
| 1 | 6 | 3.36 | 3.3272 | 3.6322 | 3.3822 | 3.16 | 3.7088 | 3.47 | 3.2825 | 3.27 | 3.15 | 3.56 | 3.41 | 3.33 | 3.4 | 3.44 | 3.544 | 3.37 | 3.374 | 1.08 | 3.0347 | 3.5915 | 3.52 | 3.4738 | 3.49 | 2.7 | 3.44 | 3.5
| Mean |  | 3.355 | 3.3051 | 3.6363 | 3.4045 | 3.2267 | 3.6837 | 3.52 | 3.3245 | 3.2717 | 3.1033 | 3.4817 | 3.5117 | 3.395 | 3.39 | 3.5333 | 3.5253 | 3.4233 | 3.353 | 1.1083 | 2.9797 | 3.5765 | 3.5017 | 3.5288 | 3.52 | 2.7583 | 3.47 | 3.4333
| Median |  | 3.36 | 3.3047 | 3.6401 | 3.4118 | 3.195 | 3.6757 | 3.535 | 3.2978 | 3.27 | 3.1 | 3.465 | 3.545 | 3.41 | 3.39 | 3.53 | 3.5435 | 3.435 | 3.352 | 1.13 | 3.0005 | 3.5753 | 3.51 | 3.5375 | 3.525 | 2.765 | 3.47 | 3.55
| Std Dev. |  | 0.042778 | 0.020422 | 0.027763 | 0.040301 | 0.11308 | 0.032692 | 0.032249 | 0.078238 | 0.013292 | 0.042269 | 0.056362 | 0.07679 | 0.071484 | 0.032249 | 0.062823 | 0.032197 | 0.044572 | 0.061592 | 0.45745 | 0.062922 | 0.013142 | 0.036009 | 0.039791 | 0.044272 | 0.061128 | 0.018974 | 0.39833
| Rel.Std.Dev. |  | 1.275% | 0.6179% | 0.7635% | 1.184% | 3.504% | 0.8875% | 0.9162% | 2.353% | 0.4063% | 1.362% | 1.619% | 2.187% | 2.106% | 0.9513% | 1.778% | 0.9133% | 1.302% | 1.837% | 41.27% | 2.112% | 0.3674% | 1.028% | 1.128% | 1.258% | 2.216% | 0.5468% | 11.6%
| PDM3 |  | -2.407% | -3.857% | 5.775% | -0.967% | -6.14% | 7.154% | 2.392% | -3.295% | -4.831% | -9.728% | 1.277% | 2.15% | -1.244% | -1.389% | 2.78% | 2.547% | -0.4197% | -2.466% | -67.76% | -13.32% | 4.036% | 1.859% | 2.647% | 2.392% | -19.76% | 0.9378% | -0.1288%
| Z-Score (Absolute) |  | 0.47 | 0.81 | 1.45 | 0.13 | 1.34 | 1.78 | 0.66 | 0.67 | 1.03 | 2.18 | 0.4 | 0.6 | 0.19 | 0.23 | 0.75 | 0.7 | 0 | 0.48 | 15.79 | 3.03 | 1.04 | 0.53 | 0.72 | 0.66 | 4.54 | 0.32 | 0.07
| Analytical results for Mn in OREAS 240b (Certified Value 0.14 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.14 | 0.1395 | 0.14854 | 0.13883 | 0.136 | 0.1549 | 0.141 | 0.1365 | 0.1272 | 0.1567 | 0.134 | 0.139 | 0.1415 | 0.137 | 0.1458 | 0.1403 | 0.1413 | 0.1397 | 0.084 | 0.1311 | 0.1385 | 0.133 | 0.12632 | 0.143 | 0.1383 | 0.1376 | 0.14 |  |  | 1
| 1 | 2 | 0.14 | 0.1407 | 0.1487 | 0.1403 | 0.13 | 0.1549 | 0.141 | 0.1384 | 0.1275 | 0.1535 | 0.136 | 0.145 | 0.1445 | 0.1365 | 0.1411 | 0.1428 | 0.141 | 0.1378 | 0.073 | 0.13026 | 0.138 | 0.1375 | 0.12794 | 0.146 | 0.1418 | 0.1362 | 0.146 |  |  | 23
| 1 | 3 | 0.14 | 0.1431 | 0.14938 | 0.14037 | 0.135 | 0.1533 | 0.141 | 0.1425 | 0.1274 | 0.1468 | 0.1325 | 0.144 | 0.141 | 0.1355 | 0.1427 | 0.1431 | 0.1418 | 0.1343 | 0.093 | 0.13438 | 0.137 | 0.136 | 0.12719 | 0.143 | 0.1428 | 0.1376 | 0.145 |  |  | 16
| 1 | 4 | 0.14 | 0.1409 | 0.14781 | 0.13826 | 0.135 | 0.1564 | 0.141 | 0.1379 | 0.1272 | 0.1413 | 0.133 | 0.145 | 0.1415 | 0.1375 | 0.1422 | 0.143 | 0.1397 | 0.1368 | 0.082 | 0.13577 | 0.1375 | 0.138 | 0.12772 | 0.145 | 0.1393 | 0.1366 | 0.139 |  |  | 0.1396
| 1 | 5 | 0.14 | 0.1379 | 0.14593 | 0.13745 | 0.145 | 0.1549 | 0.139 | 0.1343 | 0.1277 | 0.1417 | 0.139 | 0.1445 | 0.1355 | 0.138 | 0.1411 | 0.1406 | 0.1377 | 0.143 | 0.089 | 0.13446 | 0.1385 | 0.137 | 0.12578 | 0.146 | 0.136 | 0.1364 | 0.147 |  |  | 43
| 1 | 6 | 0.14 | 0.1442 | 0.14927 | 0.13791 | 0.132 | 0.1557 | 0.138 | 0.1362 | 0.1276 | 0.1614 | 0.137 | 0.1365 | 0.1385 | 0.1375 | 0.1375 | 0.1431 | 0.1376 | 0.1388 | 0.074 | 0.13686 | 0.139 | 0.1375 | 0.12546 | 0.143 | 0.139 | 0.1355 | 0.145
| Mean |  | 0.14 | 0.14105 | 0.14827 | 0.13886 | 0.1355 | 0.15502 | 0.14017 | 0.13763 | 0.12743 | 0.15023 | 0.13525 | 0.14233 | 0.14042 | 0.137 | 0.14173 | 0.14215 | 0.13985 | 0.1384 | 0.0825 | 0.1338 | 0.13808 | 0.1365 | 0.12674 | 0.14433 | 0.13953 | 0.13665 | 0.14367
| Median |  | 0.14 | 0.1408 | 0.14862 | 0.13855 | 0.135 | 0.1549 | 0.141 | 0.1372 | 0.12745 | 0.15015 | 0.135 | 0.14425 | 0.14125 | 0.13725 | 0.14165 | 0.1429 | 0.14035 | 0.1383 | 0.083 | 0.13442 | 0.13825 | 0.13725 | 0.12675 | 0.144 | 0.13915 | 0.1365 | 0.145
| Std Dev. |  | 0 | 0.0023063 | 0.0012784 | 0.0012325 | 0.0051672 | 0.0010362 | 0.0013292 | 0.0027854 | 0.00020656 | 0.0082655 | 0.0025249 | 0.003656 | 0.0030727 | 0.00089443 | 0.0026986 | 0.0013248 | 0.0018404 | 0.0029237 | 0.0079687 | 0.0026016 | 0.00073598 | 0.0018439 | 0.0010345 | 0.0015055 | 0.0024557 | 0.00082401 | 0.0033267
| Rel.Std.Dev. |  | 0% | 1.635% | 0.8622% | 0.8876% | 3.813% | 0.6684% | 0.9483% | 2.024% | 0.1621% | 5.502% | 1.867% | 2.569% | 2.188% | 0.6529% | 1.904% | 0.9319% | 1.316% | 2.112% | 9.659% | 1.944% | 0.533% | 1.351% | 0.8162% | 1.043% | 1.76% | 0.603% | 2.316%
| PDM3 |  | 0.2882% | 1.04% | 6.214% | -0.5313% | -2.935% | 11.05% | 0.4076% | -1.407% | -8.714% | 7.619% | -3.114% | 1.96% | 0.5867% | -1.861% | 1.53% | 1.828% | 0.1807% | -0.858% | -40.9% | -4.15% | -1.085% | -2.219% | -9.213% | 3.392% | -0.04611% | -2.112% | 2.915%
| Z-Score (Absolute) |  | 0.11 | 0.37 | 2.1 | 0.16 | 0.97 | 3.73 | 0.15 | 0.46 | 2.91 | 2.58 | 1.03 | 0.67 | 0.21 | 0.61 | 0.53 | 0.63 | 0.08 | 0.27 | 13.74 | 1.38 | 0.35 | 0.73 | 3.08 | 1.16 | 0 | 0.69 | 1
| Analytical results for Mo in OREAS 240b (Certified Value 2.9 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.21 | 2.9 | 3.44 | 2.7756 | 2.7162 | < 1 | 3.05 | 2.8 | 2.9 | 3 | 2.8 | 2.82 | 3.08 | 2.77 | 2.88 | 3 | 2.5 | 2.6 | 2.59 | 2.38 | 3 | 2.9201 | 2.98 | 2.776 | 3.11 | 3.46 | 2.74 | 3.05 |  | 1
| 1 | 2 | 3.32 | 3 | 3.48 | 2.7734 | 2.8289 | < 1 | 3.39 | 2.9 | 2.8 | 3 | 2.9 | 2.92 | 3.16 | 2.96 | 2.81 | 2 | 2.5 | 2.5 | 2.69 | 2.59 | 3 | 2.8584 | 3.07 | 2.7123 | 3.04 | 3.32 | 2.84 | 3.18 |  | 24
| 1 | 3 | 3.32 | 3.1 | 3.6 | 2.7932 | 2.8293 | < 1 | 3.4 | 3.1 | 3.1 | 3 | 2.7 | 2.8 | 3.04 | 2.81 | 2.75 | 3 | 2.5 | 2.5 | 2.76 | 2 | 3 | 2.9125 | 3.03 | 2.7496 | 3.03 | 3.37 | 2.88 | 2.76 |  | 16
| 1 | 4 | 3.27 | 2.9 | 2.96 | 2.8224 | 2.8169 | 3 | 3.37 | 2.9 | 2.9 | 3 | 2.6 | 2.73 | 2.92 | 2.86 | 2.88 | 3 | 2.5 | 2.5 | 2.63 | 3.6 | 3 | 2.968 | 3.07 | 2.7741 | 3.07 | 3.08 | 2.83 | 3.15 |  | 2.9033
| 1 | 5 | 3.18 | 2.8 | 3.35 | 2.699 | 2.6893 | 2 | 3.34 | 3.1 | 2.8 | 3 | 3.3 | 2.82 | 3.04 | 2.71 | 2.81 | 2 | 2.4 | 2.4 | 2.65 | 3.37 | 3 | 2.8921 | 3.01 | 2.6658 | 3.17 | 3.3 | 2.79 | 2.73 |  | 44
| 1 | 6 | 3.25 | 2.9 | 3.06 | 2.789 | 2.7367 | 3 | 3.76 | 3 | 2.8 | 3 | 3.1 | 2.82 | 2.63 | 2.75 | 2.79 | 2 | 2.5 | 2.5 | 2.72 | 2.94 | 3 | 2.9712 | 3.04 | 2.5746 | 3.13 | 3.24 | 2.87 | 2.96
| Mean |  | 3.2583 | 2.9333 | 3.315 | 2.7754 | 2.7696 | 2.6667 | 3.385 | 2.9667 | 2.8833 | 3 | 2.9 | 2.8183 | 2.9783 | 2.81 | 2.82 | 2.5 | 2.4833 | 2.5 | 2.6733 | 2.8133 | 3 | 2.9204 | 3.0333 | 2.7087 | 3.0917 | 3.295 | 2.825 | 2.9717
| Median |  | 3.26 | 2.9 | 3.395 | 2.7823 | 2.7768 | 3 | 3.38 | 2.95 | 2.85 | 3 | 2.85 | 2.82 | 3.04 | 2.79 | 2.81 | 2.5 | 2.5 | 2.5 | 2.67 | 2.765 | 3 | 2.9163 | 3.035 | 2.7309 | 3.09 | 3.31 | 2.835 | 3.005
| Std Dev. |  | 0.057067 | 0.10328 | 0.25154 | 0.041356 | 0.062786 | 0.57735 | 0.22599 | 0.12111 | 0.1169 | 0 | 0.26077 | 0.0608 | 0.18745 | 0.089666 | 0.051381 | 0.54772 | 0.040825 | 0.063246 | 0.062183 | 0.60728 | 0 | 0.043722 | 0.035024 | 0.077866 | 0.054559 | 0.12865 | 0.05244 | 0.19219
| Rel.Std.Dev. |  | 1.751% | 3.521% | 7.588% | 1.49% | 2.267% | 21.65% | 6.676% | 4.082% | 4.054% | 0% | 8.992% | 2.157% | 6.294% | 3.191% | 1.822% | 21.91% | 1.644% | 2.53% | 2.326% | 21.59% | 0% | 1.497% | 1.155% | 2.875% | 1.765% | 3.904% | 1.856% | 6.467%
| PDM3 |  | 12.23% | 1.034% | 14.18% | -4.405% | -4.608% | -8.151% | 16.59% | 2.182% | -0.6885% | 3.33% | -0.1145% | -2.927% | 2.584% | -3.214% | -2.87% | -13.89% | -14.47% | -13.89% | -7.922% | -3.1% | 3.33% | 0.5874% | 4.478% | -6.702% | 6.487% | 13.49% | -2.698% | 2.354%
| Z-Score (Absolute) |  | 2.44 | 0.28 | 2.81 | 0.77 | 0.81 | NA | 3.28 | 0.5 | 0.06 | NA | 0.06 | 0.49 | 0.58 | 0.54 | 0.48 | NA | 2.71 | 2.6 | 1.45 | 0.52 | NA | 0.19 | 0.94 | 1.22 | 1.33 | 2.68 | 0.44 | 0.53
Results from laboratories 1.07, 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 240b (Certified Value 1.87 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.9 | 1.8459 | 1.8363 | 1.7764 | 1.9102 | 1.81 | 2.033 | 1.85 | 1.7406 | 1.81 | 1.88 | 1.88 | 1.86 | 1.87 | 1.8 | 2 | 1.833 | 1.852 | 1.808 | 3.428 | 1.8458 | 1.935 | 1.87 | 2.404 | 1.89 | 2.27 | 2 | 1.76 |  | 1
| 1 | 2 | 1.92 | 1.8561 | 1.7796 | 1.7791 | 1.9104 | 1.76 | 2.062 | 1.86 | 1.7426 | 1.79 | 1.89 | 1.9 | 1.95 | 1.91 | 1.78 | 1.93 | 1.862 | 1.87 | 1.775 | 3.865 | 1.8334 | 1.92 | 1.92 | 2.384 | 1.93 | 2.16 | 1.99 | 1.86 |  | #N/A
| 1 | 3 | 1.94 | 1.8705 | 1.7533 | 1.783 | 1.9085 | 1.86 | 2.04 | 1.86 | 1.8159 | 1.77 | 1.88 | 1.87 | 1.93 | 1.85 | 1.79 | 1.96 | 1.875 | 1.879 | 1.725 | 2.535 | 1.8919 | 1.905 | 1.91 | 2.413 | 1.9 | 2.26 | 2.01 | 1.87 |  | 16
| 1 | 4 | 1.92 | 1.8449 | 1.886 | 1.7652 | 1.8668 | 1.82 | 2.047 | 1.87 | 1.7768 | 1.78 | 1.89 | 1.85 | 1.92 | 1.87 | 1.82 | 1.94 | 1.86 | 1.849 | 1.89 | 4.323 | 1.916 | 1.92 | 1.9 | 2.393 | 1.93 | 2.22 | 2 | 1.77 |  | 1.8708
| 1 | 5 | 1.87 | 1.8745 | 1.9944 | 1.7376 | 1.8562 | 1.95 | 2.07 | 1.84 | 1.7236 | 1.8 | 1.93 | 1.93 | 1.91 | 1.8 | 1.82 | 1.92 | 1.838 | 1.832 | 1.838 | 4.015 | 1.9329 | 1.925 | 1.89 | 2.342 | 1.94 | 2.19 | 1.99 | 1.91 |  | 45
| 1 | 6 | 1.92 | 1.8787 | 1.8792 | 1.7702 | 1.8664 | 1.83 | 2.025 | 1.82 | 1.725 | 1.81 | 1.93 | 1.92 | 1.86 | 1.82 | 1.8 | 1.87 | 1.864 | 1.828 | 1.785 | 3.538 | 1.9255 | 1.93 | 1.91 | 2.344 | 1.9 | 2.15 | 1.99 | 1.87
| Mean |  | 1.9117 | 1.8618 | 1.8548 | 1.7686 | 1.8864 | 1.8383 | 2.0462 | 1.85 | 1.7541 | 1.7933 | 1.9 | 1.8917 | 1.905 | 1.8533 | 1.8017 | 1.9367 | 1.8553 | 1.8517 | 1.8035 | 3.6173 | 1.8909 | 1.9225 | 1.9 | 2.38 | 1.915 | 2.2083 | 1.9967 | 1.84
| Median |  | 1.92 | 1.8633 | 1.8578 | 1.7733 | 1.8876 | 1.825 | 2.0435 | 1.855 | 1.7416 | 1.795 | 1.89 | 1.89 | 1.915 | 1.86 | 1.8 | 1.935 | 1.861 | 1.8505 | 1.7965 | 3.7015 | 1.904 | 1.9225 | 1.905 | 2.3885 | 1.915 | 2.205 | 1.995 | 1.865
| Std Dev. |  | 0.024014 | 0.014788 | 0.086461 | 0.016453 | 0.025808 | 0.063692 | 0.017198 | 0.017889 | 0.03585 | 0.01633 | 0.023664 | 0.030605 | 0.037283 | 0.039328 | 0.016021 | 0.043205 | 0.016293 | 0.020186 | 0.056585 | 0.62145 | 0.042265 | 0.010368 | 0.017889 | 0.030299 | 0.020736 | 0.050365 | 0.008165 | 0.060663
| Rel.Std.Dev. |  | 1.256% | 0.7943% | 4.662% | 0.9303% | 1.368% | 3.465% | 0.8405% | 0.9669% | 2.044% | 0.9106% | 1.245% | 1.618% | 1.957% | 2.122% | 0.8892% | 2.231% | 0.8782% | 1.09% | 3.138% | 17.18% | 2.235% | 0.5393% | 0.9415% | 1.273% | 1.083% | 2.281% | 0.4089% | 3.297%
| PDM3 |  | 2.183% | -0.4838% | -0.8566% | -5.465% | 0.8332% | -1.736% | 9.373% | -1.113% | -6.24% | -4.142% | 1.56% | 1.114% | 1.827% | -0.9346% | -3.696% | 3.52% | -0.8277% | -1.024% | -3.598% | 93.36% | 1.075% | 2.763% | 1.56% | 27.22% | 2.362% | 18.04% | 6.727% | -1.647%
| Z-Score (Absolute) |  | 0.41 | 0.48 | 0.6 | 2.14 | 0.04 | 0.9 | 2.81 | 0.69 | 2.4 | 1.7 | 0.2 | 0.05 | 0.29 | 0.63 | 1.55 | 0.86 | 0.59 | 0.66 | 1.52 | 30.81 | 0.04 | 0.6 | 0.2 | 8.76 | 0.47 | 5.7 | 1.92 | 0.87
| Analytical results for Nb in OREAS 240b (Certified Value 3.54 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.6 | 3.57 | 18.93 | 2.043 | 3.5261 | 1 | 3.5 | 3.5 | 3.3 | < 5 | 3.6 | 3.6 | 3.6 | 3.4 | 3.4 | 23 | 3.1 | 2.9 | 3.5 | 4 | 3.0771 | 3.7 | 4.8537 | 4 | 3.8 | 3.6 | 3.5 |  |  | 1
| 1 | 2 | 3.6 | 3.64 | 20.05 | 2.1298 | 3.6155 | 0.3 | 4.1 | 3.5 | 3.36 | < 5 | 3.6 | 3.7 | 3.7 | 3.7 | 3.3 | 22 | 3.3 | 3 | 3.3 | 5 | 2.9038 | 3.7 | 4.9064 | 4.1 | 3.6 | 3.5 | 3.5 |  |  | 7
| 1 | 3 | 3.5 | 3.64 | 17.23 | 2.0786 | 3.5991 | 0.4 | 3.8 | 3.6 | 3.49 | < 5 | 3.6 | 3.6 | 3.7 | 3.4 | 3.4 | 22 | 3.3 | 3 | 3.4 | 4 | 3.3177 | 3.8 | 4.8557 | 3.9 | 3.8 | 3.7 | 3.3 |  |  | 16
| 1 | 4 | 3.4 | 3.54 | 18.18 | 2.1355 | 3.5634 | 2.4 | 4 | 3.5 | 3.44 | < 5 | 3.6 | 3.6 | 3.5 | 3.6 | 3.5 | 22 | 3.3 | 2.9 | 3.5 | 5 | 3.0465 | 3.8 | 4.8403 | 4 | 3.7 | 3.7 | 3.6 |  |  | 3.5405
| 1 | 5 | 3.3 | 3.39 | 19.07 | 1.9942 | 3.5431 | 4 | 4 | 3.5 | 3.38 | < 5 | 3.6 | 3.6 | 3.5 | 4.2 | 3.3 | 22 | 3.1 | 2.8 | 3.4 | 5 | 3.0648 | 3.7 | 4.672 | 4 | 3.6 | 3.3 | 3.3 |  |  | 46
| 1 | 6 | 3.4 | 3.74 | 20.02 | 2.1092 | 3.5181 | 3.8 | 4.5 | 3.5 | 3.32 | < 5 | 3.7 | 3.7 | 3.3 | 3.4 | 3.4 | 21 | 3.4 | 2.9 | 3.4 | 4 | 3.2469 | 3.7 | 4.5431 | 4.1 | 3.6 | 3.4 | 3.5
| Mean |  | 3.4667 | 3.5867 | 18.9133 | 2.0817 | 3.5609 | 1.9833 | 3.9833 | 3.5167 | 3.3817 |  | 3.6167 | 3.6333 | 3.55 | 3.6167 | 3.3833 | 22 | 3.25 | 2.9167 | 3.4167 | 4.5 | 3.1094 | 3.7333 | 4.7785 | 4.0167 | 3.6833 | 3.5333 | 3.45
| Median |  | 3.45 | 3.605 | 19 | 2.0939 | 3.5532 | 1.7 | 4 | 3.5 | 3.37 |  | 3.6 | 3.6 | 3.55 | 3.5 | 3.4 | 22 | 3.3 | 2.9 | 3.4 | 4.5 | 3.0709 | 3.7 | 4.847 | 4 | 3.65 | 3.55 | 3.5
| Std Dev. |  | 0.12111 | 0.1186 | 1.0875 | 0.05497 | 0.039502 | 1.6642 | 0.33116 | 0.040825 | 0.072226 |  | 0.040825 | 0.05164 | 0.15166 | 0.31252 | 0.075277 | 0.63246 | 0.12247 | 0.075277 | 0.075277 | 0.54772 | 0.14936 | 0.05164 | 0.14039 | 0.075277 | 0.098319 | 0.1633 | 0.12247
| Rel.Std.Dev. |  | 3.493% | 3.307% | 5.75% | 2.641% | 1.109% | 83.91% | 8.314% | 1.161% | 2.136% |  | 1.129% | 1.421% | 4.272% | 8.641% | 2.225% | 2.875% | 3.768% | 2.581% | 2.203% | 12.17% | 4.803% | 1.383% | 2.938% | 1.874% | 2.669% | 4.622% | 3.55%
| PDM3 |  | -2.084% | 1.305% | 434.2% | -41.2% | 0.5771% | -43.98% | 12.51% | -0.6719% | -4.485% |  | 2.153% | 2.623% | 0.2696% | 2.153% | -4.438% | 521.4% | -8.204% | -17.62% | -3.496% | 27.1% | -12.17% | 5.448% | 34.97% | 13.45% | 4.036% | -0.2012% | -2.555%
| Z-Score (Absolute) |  | 0.3 | 0.24 | 69.14 | 6.53 | 0.12 | 6.97 | 2.02 | 0.07 | 0.68 | 4.65 | 0.37 | 0.45 | 0.07 | 0.37 | 0.67 | NA | 1.27 | 2.77 | 0.52 | NA | 1.91 | 0.9 | 5.6 | 2.17 | 0.67 | 0 | 0.37
Results from laboratories 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 240b (Certified Value 8.74 ppm) | [30 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 8.76 | 5.13 | 8.2661 | 8.5 | 8.3 | 8.6 | 9 | 8.8356 | [21 empty] | 1
| 1 | 2 | 9 | 5.54 | 8.2699 | 8.2 | 9.3 | 8.6 | 9 | 8.7378 | [21 empty] | 26
| 1 | 3 | 9.11 | 5.47 | 8.2719 | 8.5 | 8.8 | 9 | 8.5 | 9.2686 | [21 empty] | 16
| 1 | 4 | 9.03 | 4.97 | 8.3007 | 8.3 | 8.9 | 8.8 | 8.5 | 9.0871 | [21 empty] | 8.742
| 1 | 5 | 8.67 | 5.06 | 8.2402 | 9.3 | 9.1 | 8.7 | 9 | 9.3292 | [21 empty] | 47
| 1 | 6 | 9.28 | 5.5 | 8.2719 | 8.3 | 10.4 | 8.9 | 9 | 9.3955
| Mean |  | 8.975 | 5.2783 | 8.2701 | 8.5167 | 9.1333 | 8.7667 | 8.8333 | 9.109
| Median |  | 9.015 | 5.3 | 8.2709 | 8.4 | 9 | 8.75 | 9 | 9.1779
| Std Dev. |  | 0.22546 | 0.25262 | 0.019263 | 0.40208 | 0.70616 | 0.1633 | 0.2582 | 0.2717
| Rel.Std.Dev. |  | 2.512% | 4.786% | 0.2329% | 4.721% | 7.732% | 1.863% | 2.923% | 2.983%
| PDM3 |  | 2.665% | -39.62% | -5.398% | -2.578% | 4.476% | 0.282% | 1.045% | 4.198%
| Z-Score (Absolute) |  | 0.4 | 8.02 | 1.21 | 0.65 | 0.76 | 0.08 | 0.08 | 0.7
| Analytical results for Ni in OREAS 240b (Certified Value 94 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 93.2 | 94.5 | 90.41 | 95.14 | 100.1858 | 93 | 102 | 95 | 92.5 | 89 | 92.7 | 98.9 | 95.2 | 94.7 | 95.9 | 90 | 92 | 96.2 | 92 | 76.84 | 90.2565 | 95 | 96.8118 | 95.6 | 76.1069 | 88.8 | 106.7 | 100.3 | 93.3 | 1
| 1 | 2 | 94.3 | 94.8 | 94.46 | 96.18 | 100.847 | 93 | 116 | 95 | 93.6 | 89 | 93.4 | 101.5 | 98.6 | 97.8 | 93.1 | 87 | 94 | 92.1 | 91 | 74.69 | 90.1542 | 95 | 98.6537 | 96.5 | 83.0957 | 87.1 | 102.6 | 97.2 | 97.3 | 27
| 1 | 3 | 95.5 | 96 | 90.3 | 95.77 | 101.5169 | 92 | 107 | 96 | 96.7 | 83 | 98 | 99.3 | 98.4 | 94 | 96.5 | 88 | 93 | 94.3 | 90 | 73.42 | 90.2007 | 95 | 96.3783 | 98.5 | 79.311 | 87.5 | 103.9 | 96 | 96.4 | 16
| 1 | 4 | 93.9 | 96.5 | 90.31 | 95.41 | 100.3951 | 93 | 112 | 95 | 93.5 | 86 | 94.6 | 97.6 | 96.3 | 95.1 | 97.9 | 87 | 93 | 91.6 | 92 | 75.89 | 92.0235 | 90 | 98.2703 | 97.4 | 78.0574 | 87.8 | 99.6 | 97.6 | 92.6 | 94.3142
| 1 | 5 | 93 | 91.7 | 90.77 | 95.05 | 99.2094 | 92 | 110 | 94 | 91 | 89 | 95.3 | 100 | 98 | 91.3 | 93.3 | 87 | 92 | 89.1 | 92 | 81.56 | 91.4562 | 90 | 97.7851 | 97.1 | 75.8576 | 86 | 99.7 | 95.9 | 99.1 | 48
| 1 | 6 | 94.5 | 97.2 | 91.47 | 95.37 | 99.819 | 92 | 125 | 95 | 92.5 | 88 | 94 | 100.5 | 102 | 91.5 | 94.6 | 85 | 94 | 91.1 | 92 | 74.22 | 91.5306 | 95 | 100.7856 | 96.5 | 76.8488 | 87.6 | 107.4 | 97.7 | 96.9
| Mean |  | 94.0667 | 95.1167 | 91.2867 | 95.4867 | 100.3289 | 92.5 | 112 | 95 | 93.3 | 87.3333 | 94.6667 | 99.6333 | 98.0833 | 94.0667 | 95.2167 | 87.3333 | 93 | 92.4 | 91.5 | 76.1033 | 90.937 | 93.3333 | 98.1141 | 96.9333 | 78.2129 | 87.4667 | 103.3167 | 97.45 | 95.9333
| Median |  | 94.1 | 95.4 | 90.59 | 95.39 | 100.2905 | 92.5 | 111 | 95 | 93 | 88.5 | 94.3 | 99.65 | 98.2 | 94.35 | 95.25 | 87 | 93 | 91.85 | 92 | 75.29 | 90.8564 | 95 | 98.0277 | 96.8 | 77.4531 | 87.55 | 103.25 | 97.4 | 96.65
| Std Dev. |  | 0.91797 | 1.959 | 1.6166 | 0.42222 | 0.80224 | 0.54772 | 7.9246 | 0.63246 | 1.911 | 2.4221 | 1.8673 | 1.353 | 2.3327 | 2.4353 | 1.8894 | 1.633 | 0.89443 | 2.5044 | 0.83666 | 2.9363 | 0.82709 | 2.582 | 1.5665 | 0.98522 | 2.7172 | 0.91579 | 3.3427 | 1.5984 | 2.4937
| Rel.Std.Dev. |  | 0.9759% | 2.06% | 1.771% | 0.4422% | 0.7996% | 0.5921% | 7.076% | 0.6657% | 2.048% | 2.773% | 1.972% | 1.358% | 2.378% | 2.589% | 1.984% | 1.87% | 0.9617% | 2.71% | 0.9144% | 3.858% | 0.9095% | 2.766% | 1.597% | 1.016% | 3.474% | 1.047% | 3.235% | 1.64% | 2.599%
| PDM3 |  | -0.2624% | 0.8509% | -3.21% | 1.243% | 6.377% | -1.924% | 18.75% | 0.7272% | -1.075% | -7.402% | 0.3737% | 5.64% | 3.996% | -0.2624% | 0.9569% | -7.402% | -1.393% | -2.03% | -2.984% | -19.31% | -3.581% | -1.04% | 4.029% | 2.777% | -17.07% | -7.26% | 9.545% | 3.325% | 1.717%
| Z-Score (Absolute) |  | 0 | 0.25 | 0.67 | 0.34 | 1.52 | 0.38 | 4.35 | 0.23 | 0.19 | 1.63 | 0.15 | 1.35 | 0.97 | 0 | 0.28 | 1.63 | 0.26 | 0.4 | 0.62 | 4.36 | 0.76 | 0.18 | 0.98 | 0.7 | 3.85 | 1.6 | 2.24 | 0.82 | 0.45
| Analytical results for P in OREAS 240b (Certified Value 0.043 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.0426 | 0.0446 | 0.042516 | 0.043384 | 0.039 | 0.0506 | 0.045 | 0.0412 | 0.0434 | 0.043 | 0.043 | 0.042 | 0.044 | 0.042 | 0.0444 | 0.043 | 0.043 | 0.043 | 0.041 | 0.043 | 0.043367 | 0.045 | 0.044 | 0.0423 | 0.0417 |  |  |  |  | 1
| 1 | 2 | 0.0427 | 0.0447 | 0.042488 | 0.043573 | 0.037 | 0.0519 | 0.046 | 0.0385 | 0.0431 | 0.043 | 0.043 | 0.044 | 0.045 | 0.041 | 0.0426 | 0.044 | 0.041 | 0.043 | 0.0425 | 0.044 | 0.042984 | 0.0426 | 0.0424 | 0.0428 | 0.0428 |  |  |  |  | #N/A
| 1 | 3 | 0.043 | 0.0445 | 0.043063 | 0.043533 | 0.038 | 0.0502 | 0.046 | 0.0448 | 0.0429 | 0.042 | 0.043 | 0.044 | 0.044 | 0.04 | 0.0429 | 0.044 | 0.042 | 0.043 | 0.0415 | 0.044 | 0.043394 | 0.0425 | 0.0433 | 0.0427 | 0.0424 |  |  |  |  | 16
| 1 | 4 | 0.0426 | 0.0438 | 0.042651 | 0.043053 | 0.04 | 0.0511 | 0.046 | 0.0437 | 0.0428 | 0.043 | 0.043 | 0.044 | 0.044 | 0.043 | 0.0433 | 0.044 | 0.041 | 0.043 | 0.0425 | 0.043 | 0.042634 | 0.0426 | 0.0429 | 0.0427 | 0.0408 |  |  |  |  | 0.042884
| 1 | 5 | 0.0415 | 0.0427 | 0.042665 | 0.042321 | 0.041 | 0.0506 | 0.046 | 0.0403 | 0.0433 | 0.045 | 0.044 | 0.044 | 0.043 | 0.042 | 0.0425 | 0.043 | 0.039 | 0.043 | 0.0425 | 0.044 | 0.042758 | 0.043 | 0.0412 | 0.0424 | 0.0432 |  |  |  |  | 49
| 1 | 6 | 0.0428 | 0.0466 | 0.042085 | 0.04297 | 0.041 | 0.0528 | 0.046 | 0.0418 | 0.0435 | 0.046 | 0.045 | 0.042 | 0.043 | 0.042 | 0.0415 | 0.044 | 0.041 | 0.043 | 0.042 | 0.044 | 0.041925 | 0.0426 | 0.0421 | 0.0422 | 0.0419
| Mean |  | 0.042533 | 0.044483 | 0.042578 | 0.043139 | 0.039333 | 0.0512 | 0.045833 | 0.041717 | 0.043167 | 0.043667 | 0.0435 | 0.043333 | 0.043833 | 0.041667 | 0.042867 | 0.043667 | 0.041167 | 0.043 | 0.042 | 0.043667 | 0.042844 | 0.04305 | 0.04265 | 0.042517 | 0.042133
| Median |  | 0.04265 | 0.04455 | 0.042583 | 0.043219 | 0.0395 | 0.05085 | 0.046 | 0.0415 | 0.0432 | 0.043 | 0.043 | 0.044 | 0.044 | 0.042 | 0.04275 | 0.044 | 0.041 | 0.043 | 0.04225 | 0.044 | 0.042871 | 0.0426 | 0.04265 | 0.04255 | 0.04215
| Std Dev. |  | 0.00052789 | 0.0012797 | 0.00031721 | 0.0004708 | 0.001633 | 0.00097775 | 0.00040825 | 0.0022833 | 0.00028048 | 0.0015055 | 0.00083666 | 0.0010328 | 0.00075277 | 0.0010328 | 0.00096056 | 0.0005164 | 0.0013292 | 7.6012e-18 | 0.00063246 | 0.0005164 | 0.00054622 | 0.00097108 | 0.00097724 | 0.00024833 | 0.00085713
| Rel.Std.Dev. |  | 1.241% | 2.877% | 0.745% | 1.091% | 4.152% | 1.91% | 0.8907% | 5.473% | 0.6498% | 3.448% | 1.923% | 2.383% | 1.717% | 2.479% | 2.241% | 1.183% | 3.229% | 1.768e-14% | 1.506% | 1.183% | 1.275% | 2.256% | 2.291% | 0.5841% | 2.034%
| PDM3 |  | -0.8169% | 3.73% | -0.7128% | 0.5957% | -8.279% | 19.39% | 6.878% | -2.721% | 0.6599% | 1.826% | 1.437% | 1.049% | 2.215% | -2.838% | -0.03965% | 1.826% | -4.004% | 0.2713% | -2.061% | 1.826% | -0.09349% | 0.3879% | -0.5449% | -0.8558% | -1.75%
| Z-Score (Absolute) |  | 0.47 | 1.5 | 0.43 | 0.14 | 3.71 | 8.3 | 2.87 | 1.3 | 0.17 | 0.67 | 0.51 | 0.34 | 0.84 | 1.35 | 0.13 | 0.67 | 1.85 | 0 | 1.01 | 0.67 | 0.16 | 0.05 | 0.35 | 0.49 | 0.88
| Analytical results for Pb in OREAS 240b (Certified Value 40.5 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 40.1 | 42.3 | 35.38 | 39.9049 | 42.5318 | 36 | 44.3 | 42 | 40.6 | 44 | 37.9 | 41.2 | 42 | 39.9 | 39.8 | 37 | 34 | 39.3 | 38 | 38.51 | 39.5552 | 41 | 39.8936 | 42.2 | 163.7555 | 42 | 43.7 | 40.8 | 46.1 | 1
| 1 | 2 | 40 | 43.2 | 35.28 | 38.0919 | 41.5438 | 36 | 47 | 43 | 40.5 | 43 | 39 | 41.9 | 42.8 | 41.9 | 41.4 | 35 | 34 | 39.7 | 38 | 30.75 | 39.5082 | 41 | 39.7726 | 41.6 | 175.5043 | 41.9 | 42.4 | 39.3 | 43.7 | 28
| 1 | 3 | 41.3 | 42.8 | 32.85 | 38.4896 | 42.4362 | 35 | 46.5 | 43 | 41.6 | 43 | 36 | 40.1 | 42.6 | 40.6 | 42 | 34 | 34 | 39.5 | 38 | 32.32 | 40.0816 | 39 | 42.6927 | 42.7 | 174.3395 | 42.9 | 43.4 | 40.9 | 43 | 16
| 1 | 4 | 40.3 | 43.8 | 36.11 | 39.1783 | 43.7327 | 35 | 47.5 | 43 | 40.8 | 43 | 37.4 | 39.8 | 40.3 | 41.4 | 43.4 | 35 | 34 | 38.4 | 43 | 41.82 | 42.4222 | 40 | 41.2802 | 42.5 | 177.5951 | 41.8 | 41.9 | 42.1 | 45.7 | 40.5454
| 1 | 5 | 39.8 | 40.9 | 34.83 | 38.5284 | 42.5851 | 33 | 47.2 | 43 | 39.5 | 44 | 38.6 | 40.6 | 41.6 | 40.7 | 39.7 | 35 | 34 | 38.8 | 38 | 39.32 | 43.475 | 40 | 40.849 | 42.1 | 170.7419 | 42 | 41.7 | 39.1 | 41.5 | 50
| 1 | 6 | 40.3 | 43.6 | 35.48 | 39.0531 | 43.3232 | 34 | 52.9 | 43 | 40.1 | 43 | 40.4 | 41.7 | 41.8 | 39.7 | 39.8 | 33 | 34 | 39 | 39 | 39.01 | 41.7266 | 40 | 43.3229 | 41.7 | 163.7422 | 42.5 | 43 | 39.3 | 42.8
| Mean |  | 40.3 | 42.7667 | 34.9883 | 38.8744 | 42.6921 | 34.8333 | 47.5667 | 42.8333 | 40.5167 | 43.3333 | 38.2167 | 40.8833 | 41.85 | 40.7 | 41.0167 | 34.8333 | 34 | 39.1167 | 39 | 36.955 | 41.1281 | 40.1667 | 41.3018 | 42.1333 | 170.9464 | 42.1833 | 42.6833 | 40.25 | 43.8
| Median |  | 40.2 | 43 | 35.33 | 38.7908 | 42.5584 | 35 | 47.1 | 43 | 40.55 | 43 | 38.25 | 40.9 | 41.9 | 40.65 | 40.6 | 35 | 34 | 39.15 | 38 | 38.76 | 40.9041 | 40 | 41.0646 | 42.15 | 172.5407 | 42 | 42.7 | 40.05 | 43.35
| Std Dev. |  | 0.52536 | 1.0633 | 1.1257 | 0.64285 | 0.76188 | 1.169 | 2.8535 | 0.40825 | 0.70261 | 0.5164 | 1.4972 | 0.85654 | 0.88938 | 0.84617 | 1.5158 | 1.3292 | 0 | 0.47924 | 2 | 4.3792 | 1.6574 | 0.75277 | 1.4523 | 0.43205 | 6.0024 | 0.42622 | 0.81343 | 1.2062 | 1.7799
| Rel.Std.Dev. |  | 1.304% | 2.486% | 3.217% | 1.654% | 1.785% | 3.356% | 5.999% | 0.9531% | 1.734% | 1.192% | 3.918% | 2.095% | 2.125% | 2.079% | 3.696% | 3.816% | 0% | 1.225% | 5.128% | 11.85% | 4.03% | 1.874% | 3.516% | 1.025% | 3.511% | 1.01% | 1.906% | 2.997% | 4.064%
| PDM3 |  | -0.6052% | 5.479% | -13.71% | -4.121% | 5.295% | -14.09% | 17.32% | 5.643% | -0.07081% | 6.876% | -5.743% | 0.8335% | 3.218% | 0.3814% | 1.162% | -14.09% | -16.14% | -3.524% | -3.811% | -8.855% | 1.437% | -0.934% | 1.866% | 3.916% | 321.6% | 4.04% | 5.273% | -0.7285% | 8.027%
| Z-Score (Absolute) |  | 0.22 | 0.7 | 2.2 | 0.75 | 0.67 | 2.26 | 2.49 | 0.73 | 0.14 | 0.91 | 0.99 | 0 | 0.36 | 0.07 | 0.05 | 2.26 | 2.57 | 0.66 | 0.7 | 1.46 | 0.09 | 0.27 | 0.16 | 0.47 | 48.49 | 0.48 | 0.67 | 0.24 | 1.09
| Analytical results for Pr in OREAS 240b (Certified Value 1.91 ppm) | [30 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2 | 1.75 | 1.8339 | 1.8 | 2 | 1.92 | 2 | 1.9349 | [21 empty] | 1
| 1 | 2 | 1.97 | 1.86 | 1.782 | 1.9 | 2.2 | 1.95 | 2 | 1.8969 | [21 empty] | 29
| 1 | 3 | 2.03 | 1.85 | 1.7973 | 1.8 | 2.1 | 1.97 | 1.9 | 2.0637 | [21 empty] | 16
| 1 | 4 | 1.97 | 1.74 | 1.8186 | 1.9 | 2.1 | 1.82 | 1.9 | 1.9647 | [21 empty] | 1.9111
| 1 | 5 | 1.93 | 1.75 | 1.809 | 2 | 2.2 | 1.94 | 2 | 1.9769 | [21 empty] | 51
| 1 | 6 | 2.04 | 1.84 | 1.8309 | 1.9 | 2.4 | 1.97 | 1.9 | 2.1281
| Mean |  | 1.99 | 1.7983 | 1.8119 | 1.8833 | 2.1667 | 1.9283 | 1.95 | 1.9942
| Median |  | 1.985 | 1.795 | 1.8138 | 1.9 | 2.15 | 1.945 | 1.95 | 1.9708
| Std Dev. |  | 0.041473 | 0.057067 | 0.020025 | 0.075277 | 0.13663 | 0.056362 | 0.054772 | 0.085933
| Rel.Std.Dev. |  | 2.084% | 3.173% | 1.105% | 3.997% | 6.306% | 2.923% | 2.809% | 4.309%
| PDM3 |  | 4.128% | -5.901% | -5.19% | -1.454% | 13.37% | 0.9009% | 2.035% | 4.348%
| Z-Score (Absolute) |  | 0.62 | 1.71 | 1.55 | 0.68 | 2.77 | 0.13 | 0.13 | 0.67
| Analytical results for Rb in OREAS 240b (Certified Value 21.1 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 21.4 | 20 | 19.3423 | 19.9309 | 20.2 | 19.8 | 20.8 | 20.14 | 19.2 | 21.1 | 22.1 | 19.6 | 21 | 18.6 | 20.8 | 20.5 | 21.5 | 23.2352 | 21.8 | 21.5913 | 23.3 | 21.8 | 21.3 | 22.5 | [5 empty] | 1
| 1 | 2 | 21.7 | 20.44 | 19.3937 | 19.876 | 20.5 | 22 | 21.1 | 20.11 | 19.4 | 21.9 | 23.1 | 20.5 | 20.8 | 23.2 | 20.2 | 20.7 | 21.5 | 23.5298 | 21.2 | 22.098 | 22.5 | 21.1 | 20.9 | 21.6 | [5 empty] | 30
| 1 | 3 | 21.5 | 20.98 | 19.4369 | 19.8578 | 20.7 | 20.8 | 21.4 | 21.29 | 18.9 | 21.3 | 22 | 19.7 | 21 | 21.1 | 20.6 | 20.9 | 21 | 21.2956 | 22.4 | 22.0237 | 22.9 | 22 | 21.2 | 20.3 | [5 empty] | 16
| 1 | 4 | 21.4 | 20.94 | 18.7782 | 19.5087 | 20.6 | 21.7 | 22.2 | 20.81 | 18.4 | 21.2 | 21.2 | 19.8 | 20.9 | 20.9 | 20.1 | 19.1 | 21 | 23.4765 | 21.7 | 21.7343 | 22.8 | 21.7 | 21.8 | 23.2 | [5 empty] | 21.059
| 1 | 5 | 20 | 19.69 | 19.0713 | 19.5489 | 21.8 | 21.7 | 21.2 | 20.81 | 20.9 | 22 | 21.9 | 19.5 | 20.3 | 19.3 | 19.9 | 20.2 | 21.5 | 21.9909 | 21.7 | 21.0361 | 23.2 | 20.3 | 21.2 | 19.5 | [5 empty] | 52
| 1 | 6 | 20.6 | 20.45 | 19.3662 | 19.6087 | 20.5 | 24.1 | 21.3 | 19.69 | 20.2 | 21.6 | 21.2 | 19.2 | 20.3 | 21.9 | 19.7 | 21.1 | 20.5 | 22.7688 | 21.7 | 20.2025 | 23.4 | 20.8 | 21.4 | 21.5
| Mean |  | 21.1 | 20.4167 | 19.2314 | 19.7218 | 20.7167 | 21.6833 | 21.3333 | 20.475 | 19.5 | 21.5167 | 21.9167 | 19.7167 | 20.7167 | 20.8333 | 20.2167 | 20.4167 | 21.1667 | 22.7161 | 21.75 | 21.4476 | 23.0167 | 21.2833 | 21.3 | 21.4333
| Median |  | 21.4 | 20.445 | 19.3543 | 19.7332 | 20.55 | 21.7 | 21.25 | 20.475 | 19.3 | 21.45 | 21.95 | 19.65 | 20.85 | 21 | 20.15 | 20.6 | 21.25 | 23.002 | 21.7 | 21.6628 | 23.05 | 21.4 | 21.25 | 21.55
| Std Dev. |  | 0.65727 | 0.50883 | 0.25694 | 0.1866 | 0.55648 | 1.4331 | 0.47188 | 0.59167 | 0.90774 | 0.37639 | 0.70261 | 0.43551 | 0.33116 | 1.6825 | 0.41673 | 0.71671 | 0.40825 | 0.90071 | 0.38341 | 0.71796 | 0.34303 | 0.66156 | 0.29665 | 1.3648
| Rel.Std.Dev. |  | 3.115% | 2.492% | 1.336% | 0.9462% | 2.686% | 6.609% | 2.212% | 2.89% | 4.655% | 1.749% | 3.206% | 2.209% | 1.599% | 8.076% | 2.061% | 3.51% | 1.929% | 3.965% | 1.763% | 3.348% | 1.49% | 3.108% | 1.393% | 6.368%
| PDM3 |  | 0.1946% | -3.05% | -8.678% | -6.35% | -1.626% | 2.965% | 1.303% | -2.773% | -7.403% | 2.173% | 4.073% | -6.374% | -1.626% | -1.072% | -4% | -3.05% | 0.5112% | 7.869% | 3.281% | 1.845% | 9.296% | 1.065% | 1.144% | 1.777%
| Z-Score (Absolute) |  | 0.04 | 0.83 | 2.2 | 1.63 | 0.48 | 0.64 | 0.23 | 0.76 | 1.89 | 0.44 | 0.91 | 1.64 | 0.48 | 0.35 | 1.06 | 0.83 | 0.04 | 1.83 | 0.71 | 0.36 | 2.18 | 0.17 | 0.19 | 0.35
| Analytical results for Re in OREAS 240b (Certified Value 0.002 ppm) | [30 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.29 | 1.30 | 1.31 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.01 | 0.002 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0218 | 0.005 | 0.004 | 0.006 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.003 | 0.002 | < 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0221 | 0.006 | 0.003 | 0.006 | [10 empty] | 31
| 1 | 3 | 0.002 | < 0.01 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.003 | 0.003 | 0.004 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0224 | 0.004 | 0.004 | 0.005 | [10 empty] | 16
| 1 | 4 | < 0.002 | < 0.01 | 0.005 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.003 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0205 | 0.005 | 0.003 | 0.006 | [10 empty] | 0.0024963
| 1 | 5 | < 0.002 | < 0.01 | 0.003 | < 0.05 | 0.004 | < 0.005 | 0.003 | 0.002 | 0.003 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0215 | 0.006 | 0.003 | 0.005 | [10 empty] | 53
| 1 | 6 | < 0.002 | < 0.01 | 0.003 | < 0.05 | 0.003 | < 0.005 | 0.002 | 0.003 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0192 | 0.006 | 0.003 | 0.005
| Mean |  | 0.002 |  | 0.0028333 |  | 0.0026 |  | 0.0023333 | 0.0026667 | 0.0023333 | 0.0028 |  |  |  |  | 0.002 | 0.02125 | 0.0053333 | 0.0033333 | 0.0055
| Median |  | 0.002 |  | 0.0025 |  | 0.002 |  | 0.002 | 0.003 | 0.002 | 0.003 |  |  |  |  | 0.002 | 0.02165 | 0.0055 | 0.003 | 0.0055
| Std Dev. |  |  |  | 0.001169 |  | 0.00089443 |  | 0.0005164 | 0.0005164 | 0.0005164 | 0.00083666 |  |  |  |  | 0 | 0.0011979 | 0.0008165 | 0.0005164 | 0.00054772
| Rel.Std.Dev. |  |  |  | 41.26% |  | 34.4% |  | 22.13% | 19.36% | 22.13% | 29.88% |  |  |  |  | 0% | 5.637% | 15.31% | 15.49% | 9.959%
| PDM3 |  | -19.88% |  | 13.5% |  | 4.154% |  | -6.528% | 6.825% | -6.528% | 12.17% |  |  |  |  | -19.88% | 751.3% | 113.6% | 33.53% | 120.3%
| Z-Score (Absolute) |  | 1.35 | 1.75 | 0 | 17.94 | 0.4 | 0.27 | 0.4 | 0.13 | 0.4 | 0.27 | 2020.63 | 17.94 | 0.27 | 17.94 | 0.67 | 14.9 | 2.02 | 0.4 | 2.16
| Analytical results for S in OREAS 240b (Certified Value 0.549 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.55 | 0.55 | 0.56801 | 0.56584 | 0.52 | 0.615 | 0.549 | 0.51 | 0.56 | 0.52 | 0.55 | 0.55 | 0.55 | 0.54 | 0.56 | 0.553 | 0.5 | 0.572 | 0.52117 | 0.56 | 0.53 | 0.48853 | 0.56 | 0.57 | 0.56 | 0.54 |  |  |  | 1
| 1 | 2 | 0.55 | 0.55 | 0.56818 | 0.56597 | 0.52 | 0.6 | 0.549 | 0.52 | 0.57 | 0.53 | 0.55 | 0.57 | 0.55 | 0.53 | 0.53 | 0.565 | 0.5 | 0.574 | 0.52004 | 0.555 | 0.53 | 0.50255 | 0.57 | 0.55 | 0.56 | 0.55 |  |  |  | 16
| 1 | 3 | 0.56 | 0.55 | 0.5634 | 0.55077 | 0.52 | 0.607 | 0.552 | 0.56 | 0.56 | 0.53 | 0.54 | 0.57 | 0.54 | 0.53 | 0.54 | 0.567 | 0.5 | 0.562 | 0.53303 | 0.55 | 0.55 | 0.50336 | 0.57 | 0.56 | 0.56 | 0.55 |  |  |  | 16
| 1 | 4 | 0.55 | 0.55 | 0.57282 | 0.55342 | 0.53 | 0.611 | 0.558 | 0.53 | 0.57 | 0.54 | 0.54 | 0.57 | 0.54 | 0.54 | 0.54 | 0.568 | 0.5 | 0.541 | 0.53927 | 0.545 | 0.54 | 0.51345 | 0.56 | 0.55 | 0.56 | 0.53 |  |  |  | 0.54945
| 1 | 5 | 0.55 | 0.52 | 0.56928 | 0.5452 | 0.53 | 0.628 | 0.544 | 0.48 | 0.56 | 0.55 | 0.56 | 0.57 | 0.53 | 0.54 | 0.53 | 0.552 | 0.5 | 0.579 | 0.54447 | 0.55 | 0.54 | 0.47539 | 0.58 | 0.54 | 0.56 | 0.54 |  |  |  | 54
| 1 | 6 | 0.55 | 0.54 | 0.56758 | 0.5532 | 0.54 | 0.622 | 0.541 | 0.52 | 0.56 | 0.54 | 0.56 | 0.53 | 0.53 | 0.54 | 0.52 | 0.556 | 0.5 | 0.573 | 0.54493 | 0.55 | 0.55 | 0.50609 | 0.57 | 0.55 | 0.56 | 0.55
| Mean |  | 0.55167 | 0.54333 | 0.56821 | 0.55573 | 0.52667 | 0.61383 | 0.54883 | 0.52 | 0.56333 | 0.535 | 0.55 | 0.56 | 0.54 | 0.53667 | 0.53667 | 0.56017 | 0.5 | 0.56683 | 0.53382 | 0.55167 | 0.54 | 0.49823 | 0.56833 | 0.55333 | 0.56 | 0.54333
| Median |  | 0.55 | 0.55 | 0.56809 | 0.55331 | 0.525 | 0.613 | 0.549 | 0.52 | 0.56 | 0.535 | 0.55 | 0.57 | 0.54 | 0.54 | 0.535 | 0.5605 | 0.5 | 0.5725 | 0.53615 | 0.55 | 0.54 | 0.50295 | 0.57 | 0.55 | 0.56 | 0.545
| Std Dev. |  | 0.0040825 | 0.012111 | 0.0030313 | 0.0084189 | 0.008165 | 0.010147 | 0.0059805 | 0.026077 | 0.005164 | 0.010488 | 0.0089443 | 0.016733 | 0.0089443 | 0.005164 | 0.013663 | 0.0073052 | 0 | 0.013819 | 0.01111 | 0.005164 | 0.0089443 | 0.013815 | 0.0075277 | 0.010328 | 0 | 0.008165
| Rel.Std.Dev. |  | 0.74% | 2.229% | 0.5335% | 1.515% | 1.55% | 1.653% | 1.09% | 5.015% | 0.9167% | 1.96% | 1.626% | 2.988% | 1.656% | 0.9622% | 2.546% | 1.304% | 0% | 2.438% | 2.081% | 0.9361% | 1.656% | 2.773% | 1.325% | 1.866% | 0% | 1.503%
| PDM3 |  | 0.4028% | -1.114% | 3.414% | 1.143% | -4.147% | 11.72% | -0.1129% | -5.361% | 2.526% | -2.631% | 0.09946% | 1.919% | -1.721% | -2.327% | -2.327% | 1.95% | -9% | 3.163% | -2.845% | 0.4028% | -1.721% | -9.323% | 3.436% | 0.7061% | 1.919% | -1.114%
| Z-Score (Absolute) |  | 0.13 | 0.35 | 1.08 | 0.36 | 1.31 | 3.7 | 0.03 | 1.69 | 0.8 | 0.83 | 0.03 | 0.61 | 0.54 | 0.73 | 0.73 | 0.62 | 2.84 | 1 | 0.9 | 0.13 | 0.54 | 2.94 | 1.09 | 0.22 | 0.61 | 0.35
| Analytical results for Sb in OREAS 240b (Certified Value 2.14 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.41 | 2.16 | 2.414 | 0.3 | 2 | 2.15 | < 5 | 2.4 | 2.07 | 2.16 | 2.01 | 2.23 | < 5 | 1.8 | 1.9 | 2.05 | 10.78 | 2 | 1.7882 | 2.17 | 2.1712 | 2.16 | 2.33 | 2.2 | 2.33 |  |  |  |  | 1
| 1 | 2 | 2.49 | 2.3 | 2.2649 | 0.4 | 2.1 | 2.16 | < 5 | 2.5 | 2.1 | 2.24 | 2.13 | 2.12 | < 5 | 2 | 1.9 | 2.05 | 11.17 | 2 | 1.7004 | 2.08 | 2.1651 | 2.09 | 2.28 | 2.1 | 2.27 |  |  |  |  | 32
| 1 | 3 | 2.49 | 2.22 | 2.2105 | 0.6 | 2 | 2.25 | < 5 | 2.3 | 2.03 | 2.2 | 2.07 | 2.23 | < 5 | 2 | 1.9 | 1.98 | 7.12 | 2 | 1.8761 | 2.23 | 2.178 | 2.16 | 2.34 | 2.2 | 2.16 |  |  |  |  | 16
| 1 | 4 | 2.39 | 2.25 | 2.2062 | 1.3 | 2 | 2.18 | < 5 | 2.4 | 2.07 | 2.12 | 2.15 | 2.27 | < 5 | 2 | 1.9 | 2.06 | 12.14 | 2 | 1.7626 | 2.23 | 2.1202 | 2.13 | 2.38 | 2.2 | 2.47 |  |  |  |  | 2.1393
| 1 | 5 | 2.37 | 2.11 | 2.2135 | 2.2 | 2 | 2.13 | < 5 | 2.4 | 2.03 | 2.11 | 2.06 | 2.11 | < 5 | 1.8 | 1.9 | 1.97 | 10.86 | 2 | 1.7658 | 2.35 | 1.9987 | 2.21 | 2.17 | 2 | 2.17 |  |  |  |  | 55
| 1 | 6 | 2.44 | 2.16 | 2.1988 | 1.9 | 2 | 2.18 | < 5 | 2.4 | 2.13 | 1.94 | 2.08 | 2.14 | < 5 | 2.1 | 1.9 | 2.08 | 9.18 | 2 | 1.887 | 2.25 | 1.9609 | 2.14 | 2.34 | 2 | 2.3
| Mean |  | 2.4317 | 2.2 | 2.2513 | 1.1167 | 2.0167 | 2.175 |  | 2.4 | 2.0717 | 2.1283 | 2.0833 | 2.1833 |  | 1.95 | 1.9 | 2.0317 | 10.2083 | 2 | 1.7967 | 2.2183 | 2.099 | 2.1483 | 2.3067 | 2.1167 | 2.2833
| Median |  | 2.425 | 2.19 | 2.212 | 0.95 | 2 | 2.17 |  | 2.4 | 2.07 | 2.14 | 2.075 | 2.185 |  | 2 | 1.9 | 2.05 | 10.82 | 2 | 1.777 | 2.23 | 2.1426 | 2.15 | 2.335 | 2.15 | 2.285
| Std Dev. |  | 0.050761 | 0.06957 | 0.083094 | 0.8085 | 0.040825 | 0.041352 |  | 0.063246 | 0.0392 | 0.10439 | 0.050465 | 0.06802 |  | 0.12247 | 2.4324e-16 | 0.04535 | 1.789 | 0 | 0.072009 | 0.089536 | 0.095293 | 0.039707 | 0.074207 | 0.098319 | 0.11448
| Rel.Std.Dev. |  | 2.087% | 3.162% | 3.691% | 72.4% | 2.024% | 1.901% |  | 2.635% | 1.892% | 4.905% | 2.422% | 3.115% |  | 6.281% | 1.28e-14% | 2.232% | 17.53% | 0% | 4.008% | 4.036% | 4.54% | 1.848% | 3.217% | 4.645% | 5.014%
| PDM3 |  | 13.66% | 2.835% | 5.234% | -47.8% | -5.734% | 1.667% |  | 12.18% | -3.163% | -0.5145% | -2.618% | 2.056% |  | -8.85% | -11.19% | -5.033% | 377.2% | -6.513% | -16.02% | 3.692% | -1.885% | 0.4204% | 7.821% | -1.06% | 6.731%
| Z-Score (Absolute) |  | 1.45 | 0.21 | 0.48 | 5.6 | 0.78 | 0.07 | 1.81 | 1.28 | 0.48 | 0.18 | 0.42 | 0.12 | 1.81 | 1.13 | 1.4 | 0.7 | 43.12 | NA | 1.96 | 0.3 | 0.34 | 0.07 | 0.78 | 0.24 | 0.65
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 240b (Certified Value 33.3 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 35.6 | 33 | 28.01 | 32.3408 | 33.3731 | 33 | 33.1 | 32 | 32.7 | 33 | 31.1 | 34.5 | 34.1 | 32 | 32.8 | 33 | 32 | 32 | 33.7 | 37 | 34.7649 | 35.1 | 31.5141 | 34.8 | 35 | 33 | 36.9 |  |  | 1
| 1 | 2 | 36.4 | 32.5 | 29.32 | 32.696 | 33.576 | 33 | 38.7 | 32 | 32.8 | 32 | 31.4 | 35.2 | 35.4 | 34.3 | 31.5 | 32 | 33 | 32 | 33.7 | 35 | 34.9298 | 35.1 | 31.5534 | 33.9 | 33.8 | 31.7 | 35.6 |  |  | 33
| 1 | 3 | 36.2 | 33.9 | 29.36 | 32.6768 | 33.7096 | 33 | 35.1 | 32 | 33.7 | 32 | 32.6 | 33.7 | 35 | 32.4 | 32.3 | 32 | 33 | 32 | 33 | 35 | 33.9781 | 35.2 | 31.5837 | 33.8 | 35 | 32.6 | 33.1 |  |  | 16
| 1 | 4 | 36.3 | 33.3 | 27.13 | 32.312 | 33.0106 | 33 | 36.4 | 32 | 33.4 | 32 | 31.2 | 34 | 34.1 | 32.3 | 33.6 | 32 | 33 | 31 | 32.3 | 35 | 34.9974 | 34.9 | 31.5839 | 34.2 | 34.7 | 33.6 | 37.6 |  |  | 33.328
| 1 | 5 | 34.7 | 32 | 27.62 | 32.928 | 32.8034 | 33 | 36.6 | 32 | 31.8 | 33 | 33.2 | 34 | 33.9 | 32.8 | 32 | 32 | 32 | 31 | 33.9 | 36 | 33.9588 | 34.8 | 31.3449 | 33.2 | 33.3 | 32.7 | 32.2 |  |  | 56
| 1 | 6 | 35.3 | 33.6 | 28.83 | 32.6864 | 33.0369 | 32 | 41.7 | 32 | 32.7 | 33 | 33.4 | 34.2 | 37.1 | 31.6 | 32.9 | 31 | 33 | 31 | 33.7 | 34 | 34.241 | 35.5 | 30.9858 | 34.2 | 33.5 | 32.8 | 35.8
| Mean |  | 35.75 | 33.05 | 28.3783 | 32.6067 | 33.2516 | 32.8333 | 36.9333 | 32 | 32.85 | 32.5 | 32.15 | 34.2667 | 34.9333 | 32.5667 | 32.5167 | 32 | 32.6667 | 31.5 | 33.3833 | 35.3333 | 34.4783 | 35.1 | 31.4276 | 34.0167 | 34.2167 | 32.7333 | 35.2
| Median |  | 35.9 | 33.15 | 28.42 | 32.6816 | 33.205 | 33 | 36.5 | 32 | 32.75 | 32.5 | 32 | 34.1 | 34.55 | 32.35 | 32.55 | 32 | 33 | 31.5 | 33.7 | 35 | 34.5029 | 35.1 | 31.5337 | 34.05 | 34.25 | 32.75 | 35.7
| Std Dev. |  | 0.67157 | 0.7064 | 0.92989 | 0.23665 | 0.35632 | 0.40825 | 2.9764 | 0 | 0.65955 | 0.54772 | 1.0426 | 0.52789 | 1.2144 | 0.9395 | 0.7414 | 0.63246 | 0.5164 | 0.54772 | 0.61455 | 1.0328 | 0.47578 | 0.24495 | 0.23415 | 0.53072 | 0.77309 | 0.6186 | 2.1251
| Rel.Std.Dev. |  | 1.879% | 2.137% | 3.277% | 0.7258% | 1.072% | 1.243% | 8.059% | 0% | 2.008% | 1.685% | 3.243% | 1.541% | 3.476% | 2.885% | 2.28% | 1.976% | 1.581% | 1.739% | 1.841% | 2.923% | 1.38% | 0.6979% | 0.745% | 1.56% | 2.259% | 1.89% | 6.037%
| PDM3 |  | 7.267% | -0.8342% | -14.85% | -2.164% | -0.2294% | -1.484% | 10.82% | -3.985% | -1.434% | -2.485% | -3.535% | 2.816% | 4.817% | -2.284% | -2.434% | -3.985% | -1.984% | -5.485% | 0.1659% | 6.017% | 3.451% | 5.317% | -5.702% | 2.066% | 2.666% | -1.784% | 5.617%
| Z-Score (Absolute) |  | 2.3 | 0.16 | 3.55 | 0.19 | 0.32 | 0.01 | 3.24 | 0.67 | 0 | 0.28 | 0.56 | 1.12 | 1.65 | 0.22 | 0.26 | 0.67 | 0.15 | 1.07 | 0.42 | 1.97 | 1.29 | 1.78 | 1.13 | 0.93 | 1.08 | 0.09 | 1.86
| Analytical results for Se in OREAS 240b (Indicative Value 1.05 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 2 | 0.7 | < 1.5 | 0.5 | < 1 | 0.7 | < 5 | 2 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 6.79 | < 5 | 1 | 0.6003 | 2.5 | < 2 | < 1 | 0.6 | [6 empty] | 1
| 1 | 2 | < 2 | 0.6 | < 1.5 | 0.4 | < 1 | 0.8 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 3.71 | < 5 | 1 | 0.6473 | 2.2 | < 2 | < 1 | 0.6 | [6 empty] | 7
| 1 | 3 | < 2 | 0.6 | < 1.5 | 0.5 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | 1 | 1 | 6.46 | < 5 | 1 | 0.6412 | 1.6 | < 2 | < 1 | 0.5 | [6 empty] | 16
| 1 | 4 | < 2 | 0.5 | < 1.5 | 0.6 | 1 | 0.6 | < 5 | 2 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 8.08 | < 5 | 1 | 0.6382 | 1.9 | < 2 | < 1 | 0.6 | [6 empty] | 1.049
| 1 | 5 | < 2 | 0.7 | < 1.5 | 0.7 | < 1 | 0.6 | < 5 | 2 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 5.88 | < 5 | 1 | 0.6261 | 1.8 | < 2 | < 1 | 0.5 | [6 empty] | 9
| 1 | 6 | < 2 | 0.6 | < 1.5 | 0.6 | 1 | 0.5 | < 5 | 2 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 7.62 | < 5 | 1 | 0.6237 | 2.2 | < 2 | < 1 | 0.5
| Mean |  |  | 0.61667 |  | 0.55 | 1 | 0.64 |  | 1.6667 | 1 | 1 | 2 |  |  | 1 | 1 | 6.4233 |  | 1 | 0.62947 | 2.0333 |  |  | 0.55
| Median |  |  | 0.6 |  | 0.55 | 1 | 0.6 |  | 2 | 1 | 1 | 2 |  |  | 1 | 1 | 6.625 |  | 1 | 0.63215 | 2.05 |  |  | 0.55
| Std Dev. |  |  | 0.075277 |  | 0.10488 | 0 | 0.11402 |  | 0.5164 | 0 | 0 | 0 |  |  |  | 0 | 1.5472 |  | 0 | 0.016895 | 0.3266 |  |  | 0.054772
| Rel.Std.Dev. |  |  | 12.21% |  | 19.07% | 0% | 17.82% |  | 30.98% | 0% | 0% | 0% |  |  |  | 0% | 24.09% |  | 0% | 2.684% | 16.06% |  |  | 9.959%
| PDM3 |  |  | -41.21% |  | -47.57% | -4.672% | -38.99% |  | 58.88% | -4.672% | -4.672% | 90.66% |  |  | -4.672% | -4.672% | 512.3% |  | -4.672% | -39.99% | 93.83% |  |  | -47.57%
| Z-Score (Absolute) |  | 0 | 0.62 | 0.4 | 0.73 | 0.54 | 0.69 | 2.43 | 1.08 | 0 | 0 | 1.62 | 2.43 | 0 | 0.67 | 0 | 8.78 | 2.43 | 0 | 0.6 | 1.67 | 0 | 0.81 | 0.73
| Analytical results for Sm in OREAS 240b (Certified Value 2.42 ppm) | [30 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.46 | 1.96 | 2.2828 | 2.4 | 2.2 | 2.52 | 2.4 | 2.4406 | [21 empty] | 1
| 1 | 2 | 2.54 | 2.1 | 2.2661 | 2.4 | 2.5 | 2.5 | 2.2 | 2.3947 | [21 empty] | 18
| 1 | 3 | 2.57 | 2.09 | 2.2829 | 2.5 | 2.3 | 2.61 | 2.5 | 2.6035 | [21 empty] | 16
| 1 | 4 | 2.47 | 1.94 | 2.2846 | 2.5 | 2.4 | 2.39 | 2.3 | 2.4229 | [21 empty] | 2.4241
| 1 | 5 | 2.36 | 1.97 | 2.2136 | 2.8 | 2.4 | 2.58 | 2.2 | 2.5271 | [21 empty] | 57
| 1 | 6 | 2.5 | 2.08 | 2.222 | 2.5 | 2.7 | 2.64 | 2.3 | 2.5693
| Mean |  | 2.4833 | 2.0233 | 2.2587 | 2.5167 | 2.4167 | 2.54 | 2.3167 | 2.493
| Median |  | 2.485 | 2.025 | 2.2745 | 2.5 | 2.4 | 2.55 | 2.3 | 2.4839
| Std Dev. |  | 0.073394 | 0.073937 | 0.032471 | 0.1472 | 0.17224 | 0.090554 | 0.1169 | 0.085472
| Rel.Std.Dev. |  | 2.955% | 3.654% | 1.438% | 5.849% | 7.127% | 3.565% | 5.046% | 3.428%
| PDM3 |  | 2.446% | -16.53% | -6.822% | 3.821% | -0.3047% | 4.783% | -4.43% | 2.845%
| Z-Score (Absolute) |  | 0.29 | 3.67 | 1.65 | 0.57 | 0.29 | 0.77 | 1.15 | 0.37
| Analytical results for Sn in OREAS 240b (Certified Value 1.32 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.6 | 1.3 | 6.61 | 1.3001 | 1 | 1.31 | 1.7 | 1.3 | < 10 | 1.4 | 1.2 | 1.3 | 1.2 | 1.3 | < 10 | 1.2 | 1.2 | 1.3 | 1.71 | 1.5 | 1.1265 | 1.2 | 0.9543 | 1.5 | 1.5 | 1.4 | 2.2 |  |  | 1
| 1 | 2 | 1.7 | 1.3 | 3.88 | 1.353 | 1 | 1.54 | 1.8 | 1.3 | < 10 | 1.4 | 1.2 | 1.4 | 1.3 | 1.3 | < 10 | 1.2 | 1.1 | 1.3 | 2.91 | 1.5 | 1.0643 | 1.4 | 0.9539 | 1.3 | 1.4 | 1.3 | 2.1 |  |  | 19
| 1 | 3 | 1.7 | 1.3 | 4.49 | 1.3133 | 1 | 1.39 | 1.8 | 1.3 | < 10 | 1.4 | 1.2 | 1.3 | 1.3 | 1.3 | < 10 | 1.2 | 1.4 | 1.3 | 1.81 | 1.5 | 1.2076 | 1.3 | 0.9751 | 1.4 | 1.5 | 1.4 | 1.9 |  |  | 16
| 1 | 4 | 1.7 | 1.4 | 2.29 | 1.3213 | 1 | 1.47 | 1.8 | 1.3 | < 10 | 1.5 | 1.2 | 1.3 | 1.2 | 1.4 | < 10 | 1.2 | 1.1 | 1.3 | 2.38 | 1.5 | 1.113 | 1.4 | 0.9657 | 1.3 | 1.4 | 1.4 | 2.1 |  |  | 1.3165
| 1 | 5 | 1.6 | 1.3 | 2.42 | 1.3132 | 2 | 1.41 | 1.8 | 1.3 | < 10 | 1.4 | 1.2 | 1.3 | 1.4 | 1.3 | < 10 | 1.2 | 1.1 | 1.3 | 1.26 | 1.5 | 1.1239 | 1.3 | 0.8854 | 1.4 | 1.3 | 1.3 | 1.8 |  |  | 58
| 1 | 6 | 1.7 | 1.3 | 2.92 | 1.3051 | 1 | 1.68 | 1.8 | 1.3 | < 10 | 1.4 | 1.3 | 1.2 | 1.2 | 1.2 | < 10 | 1.3 | 1.4 | 1.3 | 1.8 | 1.5 | 1.1959 | 1.3 | 0.8228 | 1.4 | 1.3 | 1.3 | 2
| Mean |  | 1.6667 | 1.3167 | 3.7683 | 1.3177 | 1.1667 | 1.4667 | 1.7833 | 1.3 |  | 1.4167 | 1.2167 | 1.3 | 1.2667 | 1.3 |  | 1.2167 | 1.2167 | 1.3 | 1.9783 | 1.5 | 1.1385 | 1.3167 | 0.9262 | 1.3833 | 1.4 | 1.35 | 2.0167
| Median |  | 1.7 | 1.3 | 3.4 | 1.3132 | 1 | 1.44 | 1.8 | 1.3 |  | 1.4 | 1.2 | 1.3 | 1.25 | 1.3 |  | 1.2 | 1.15 | 1.3 | 1.805 | 1.5 | 1.1252 | 1.3 | 0.9541 | 1.4 | 1.4 | 1.35 | 2.05
| Std Dev. |  | 0.05164 | 0.040825 | 1.6329 | 0.018787 | 0.40825 | 0.13003 | 0.040825 | 0 |  | 0.040825 | 0.040825 | 0.063246 | 0.08165 | 0.063246 |  | 0.040825 | 0.1472 | 0 | 0.57922 | 0 | 0.054008 | 0.075277 | 0.059774 | 0.075277 | 0.089443 | 0.054772 | 0.1472
| Rel.Std.Dev. |  | 3.098% | 3.101% | 43.33% | 1.426% | 34.99% | 8.865% | 2.289% | 0% |  | 2.882% | 3.355% | 4.865% | 6.446% | 4.865% |  | 3.355% | 12.1% | 0% | 29.28% | 0% | 4.744% | 5.717% | 6.454% | 5.442% | 6.389% | 4.057% | 7.299%
| PDM3 |  | 26.6% | 0.01517% | 186.2% | 0.09078% | -11.38% | 11.41% | 35.46% | -1.251% |  | 7.611% | -7.581% | -1.251% | -3.783% | -1.251% |  | -7.581% | -7.581% | -1.251% | 50.28% | 13.94% | -13.52% | 0.01517% | -29.65% | 5.079% | 6.345% | 2.547% | 53.19%
| Z-Score (Absolute) |  | 1.92 | 0.1 | 14.01 | 0.09 | NA | 0.76 | 2.59 | 0.19 | 21.1 | 0.48 | 0.67 | 0.19 | 0.39 | 0.19 | 21.1 | 0.67 | 0.67 | 0.19 | 3.71 | 0.96 | 1.12 | 0.1 | 2.35 | 0.28 | 0.38 | 0.09 | 3.93
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 240b (Certified Value 196 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 177.8 | 187.19 | 201 | 198.646 | 201.0666 | 185 | 199 | 198 | 187.47 | 184 | 186 | 199.5 | 204 | 198 | 200 | 210 | 188 | 182 | 193 | 140.8 | 184.0625 | 190 | 197.3801 | 200 | 212.7775 | 195 | 211.6 | 198.8 | 223 | 1
| 1 | 2 | 189.8 | 191.43 | 204 | 198.6068 | 200.8668 | 177 | 201 | 200 | 189.52 | 185 | 185 | 204 | 209 | 208 | 198 | 204 | 193.3 | 180 | 192 | 134.81 | 183.692 | 190 | 199.0394 | 204 | 212.5896 | 199 | 205.2 | 199.5 | 214 | 20
| 1 | 3 | 188.4 | 193.91 | 200 | 202.5562 | 203.8401 | 183 | 198 | 201 | 195.48 | 185 | 193 | 199 | 208 | 196.5 | 199 | 208 | 191.3 | 183 | 186 | 140.39 | 188.3565 | 188 | 194.1854 | 206 | 215.0122 | 195 | 218.1 | 199.1 | 199 | 16
| 1 | 4 | 183.8 | 190.61 | 194 | 197.764 | 200.0222 | 187 | 201 | 203 | 188.6 | 183 | 191 | 198.5 | 208 | 199.5 | 203 | 207 | 191.6 | 179 | 191 | 134.32 | 188.518 | 189 | 200.9347 | 204 | 214.8374 | 199 | 203.1 | 199.2 | 226 | 196.3146
| 1 | 5 | 179.4 | 191.72 | 194 | 195.9412 | 199.7001 | 195 | 202 | 201 | 181.4 | 189 | 194 | 205 | 209 | 191.5 | 203 | 205 | 187.9 | 177 | 199 | 155.47 | 190.171 | 189 | 189.7399 | 203 | 196.0413 | 200 | 200 | 197.1 | 192 | 59
| 1 | 6 | 184.9 | 192.55 | 194 | 199.5574 | 199.3054 | 180 | 199 | 199 | 186.03 | 186 | 195 | 206 | 197.5 | 194 | 201 | 198 | 191.4 | 180 | 193 | 131.59 | 191.2255 | 191 | 198.1486 | 204 | 187.9245 | 196 | 202.1 | 198.2 | 214
| Mean |  | 184.0167 | 191.235 | 197.8333 | 198.8453 | 200.8002 | 184.5 | 200 | 200.3333 | 188.0833 | 185.3333 | 190.6667 | 202 | 205.9167 | 197.9167 | 200.6667 | 205.3333 | 190.5833 | 180.1667 | 192.3333 | 139.5633 | 187.6709 | 189.5 | 196.5714 | 203.5 | 206.5304 | 197.3333 | 206.6833 | 198.65 | 211.3333
| Median |  | 184.35 | 191.575 | 197 | 198.6264 | 200.4445 | 184 | 200 | 200.5 | 188.035 | 185 | 192 | 201.75 | 208 | 197.25 | 200.5 | 206 | 191.35 | 180 | 192.5 | 137.6 | 188.4373 | 189.5 | 197.7643 | 204 | 212.6835 | 197.5 | 204.15 | 198.95 | 214
| Std Dev. |  | 4.7634 | 2.2764 | 4.4008 | 2.1899 | 1.635 | 6.253 | 1.5492 | 1.7512 | 4.6109 | 2.0656 | 4.2269 | 3.3615 | 4.5212 | 5.7045 | 2.0656 | 4.1793 | 2.1665 | 2.137 | 4.1793 | 8.5895 | 3.1281 | 1.0488 | 4.0158 | 1.9748 | 11.6008 | 2.2509 | 6.859 | 0.87807 | 13.3517
| Rel.Std.Dev. |  | 2.589% | 1.19% | 2.224% | 1.101% | 0.8142% | 3.389% | 0.7746% | 0.8741% | 2.452% | 1.115% | 2.217% | 1.664% | 2.196% | 2.882% | 1.029% | 2.035% | 1.137% | 1.186% | 2.173% | 6.155% | 1.667% | 0.5535% | 2.043% | 0.9704% | 5.617% | 1.141% | 3.319% | 0.442% | 6.318%
| PDM3 |  | -6.264% | -2.587% | 0.7736% | 1.289% | 2.285% | -6.018% | 1.877% | 2.047% | -4.193% | -5.594% | -2.877% | 2.896% | 4.891% | 0.8161% | 2.217% | 4.594% | -2.919% | -8.226% | -2.028% | -28.91% | -4.403% | -3.471% | 0.1308% | 3.66% | 5.204% | 0.5189% | 5.282% | 1.19% | 7.65%
| Z-Score (Absolute) |  | 1.3 | 0.62 | 0 | 0.1 | 0.28 | 1.25 | 0.2 | 0.24 | 0.92 | 1.18 | 0.67 | 0.39 | 0.76 | 0.01 | 0.27 | 0.71 | 0.68 | 1.66 | 0.52 | 5.48 | 0.96 | 0.78 | 0.12 | 0.53 | 0.82 | 0.05 | 0.83 | 0.08 | 1.27
| Analytical results for Ta in OREAS 240b (Certified Value 0.24 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.15 | 0.24 | 0.33638 | < 0.1 | 0.19 | 0.4 | 0.22 | 11 | 0.3 | 0.22 | 0.24 | 0.21 | 0.28 | 0.19 | 0.2 | 0.22 | < 0.5 | 0.19404 | 0.25 | 0.3491 | 0.46 | 0.76 | 0.31 | 0.26 | [5 empty] | 1
| 1 | 2 | 0.11 | 0.26 | 0.33348 | < 0.1 | 0.22 | 0.37 | 0.23 | 11 | 0.4 | 0.23 | 0.25 | 0.24 | 0.27 | 0.21 | 0.2 | 0.22 | < 0.5 | 0.19432 | 0.24 | 0.3661 | 0.51 | 0.66 | 0.3 | 0.24 | [5 empty] | 21
| 1 | 3 | 0.07 | 0.25 | 0.32623 | < 0.1 | 0.2 | 0.36 | 0.23 | 11 | 0.4 | 0.23 | 0.25 | 0.22 | 0.28 | 0.2 | 0.2 | 0.23 | < 0.5 | 0.20619 | 0.24 | 0.3233 | 0.43 | 0.8 | 0.28 | 0.23 | [5 empty] | 16
| 1 | 4 | 0.07 | 0.25 | 0.32899 | 0.2 | 0.21 | 0.37 | 0.23 | 11 | 0.3 | 0.21 | 0.24 | 0.22 | 0.27 | 0.21 | 0.2 | 0.23 | < 0.5 | 0.2034 | 0.24 | 0.3602 | 0.44 | 0.81 | 0.28 | 0.25 | [5 empty] | 0.24205
| 1 | 5 | 0.07 | 0.23 | 0.331 | 0.3 | 0.26 | 0.38 | 0.22 | 11 | 0.3 | 0.23 | 0.25 | 0.23 | 0.26 | 0.19 | 0.2 | 0.23 | < 0.5 | 0.19868 | 0.25 | 0.3408 | 0.44 | 0.7 | 0.26 | 0.25 | [5 empty] | 60
| 1 | 6 | 0.09 | 0.27 | 0.31439 | 0.2 | 0.26 | 0.35 | 0.22 | 11 | 0.4 | 0.22 | 0.23 | 0.22 | 0.25 | 0.21 | 0.2 | 0.23 | < 0.5 | 0.20526 | 0.24 | 0.3132 | 0.45 | 0.7 | 0.29 | 0.24
| Mean |  | 0.093333 | 0.25 | 0.32841 | 0.23333 | 0.22333 | 0.37167 | 0.225 | 11 | 0.35 | 0.22333 | 0.24333 | 0.22333 | 0.26833 | 0.20167 | 0.2 | 0.22667 |  | 0.20032 | 0.24333 | 0.34212 | 0.455 | 0.73833 | 0.28667 | 0.245
| Median |  | 0.08 | 0.25 | 0.33 | 0.2 | 0.215 | 0.37 | 0.225 | 11 | 0.35 | 0.225 | 0.245 | 0.22 | 0.27 | 0.205 | 0.2 | 0.23 |  | 0.20104 | 0.24 | 0.34495 | 0.445 | 0.73 | 0.285 | 0.245
| Std Dev. |  | 0.032042 | 0.014142 | 0.0077137 | 0.057735 | 0.030111 | 0.017224 | 0.0054772 | 0 | 0.054772 | 0.008165 | 0.008165 | 0.010328 | 0.01169 | 0.0098319 | 3.0405e-17 | 0.005164 |  | 0.0054125 | 0.005164 | 0.020702 | 0.02881 | 0.0608 | 0.017512 | 0.010488
| Rel.Std.Dev. |  | 34.33% | 5.657% | 2.349% | 24.74% | 13.48% | 4.634% | 2.434% | 0% | 15.65% | 3.656% | 3.355% | 4.624% | 4.357% | 4.875% | 1.52e-14% | 2.278% |  | 2.702% | 2.122% | 6.051% | 6.332% | 8.235% | 6.109% | 4.281%
| PDM3 |  | -61.44% | 3.284% | 35.68% | -3.602% | -7.733% | 53.55% | -7.045% | 4444% | 44.6% | -7.733% | 0.5294% | -7.733% | 10.86% | -16.68% | -17.37% | -6.356% |  | -17.24% | 0.5294% | 41.34% | 87.98% | 205% | 18.43% | 1.218%
| Z-Score (Absolute) |  | 4.52 | 0.17 | 2.52 | NA | 0.62 | 3.82 | 0.57 | NA | NA | 0.62 | 0.02 | 0.62 | 0.72 | 1.27 | NA | 0.52 | 0.17 | 1.31 | 0.02 | 2.94 | 6.32 | 14.81 | 1.27 | 0.02
Results from laboratories 1.07, 1.13 and 1.23 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 240b (Certified Value 0.5 ppm) | [30 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 0.58 | 0.51 | 0.46 | 0.47307 | 0.5 | 0.44 | 0.5 | 0.45 | 2.9568 | 0.54 | [19 empty] | 1
| 1 | 2 | 0.58 | 0.52 | 0.49 | 0.47679 | 0.5 | 0.5 | 0.5 | 0.5 | 2.8701 | 0.53 | [19 empty] | 22
| 1 | 3 | 0.59 | 0.54 | 0.49 | 0.48712 | 0.5 | 0.46 | 0.51 | 0.5 | 3.1488 | 0.53 | [19 empty] | 16
| 1 | 4 | 0.57 | 0.53 | 0.46 | 0.48831 | 0.5 | 0.47 | 0.51 | 0.45 | 3.014 | 0.51 | [19 empty] | 0.50369
| 1 | 5 | 0.56 | 0.51 | 0.47 | 0.46569 | 0.6 | 0.49 | 0.51 | 0.45 | 3.0628 | 0.48 | [19 empty] | 61
| 1 | 6 | 0.58 | 0.53 | 0.49 | 0.47623 | 0.6 | 0.55 | 0.52 | 0.45 | 3.244 | 0.5
| Mean |  | 0.57667 | 0.52333 | 0.47667 | 0.47787 | 0.53333 | 0.485 | 0.50833 | 0.46667 | 3.0494 | 0.515
| Median |  | 0.58 | 0.525 | 0.48 | 0.47651 | 0.5 | 0.48 | 0.51 | 0.45 | 3.0384 | 0.52
| Std Dev. |  | 0.010328 | 0.012111 | 0.015055 | 0.0086008 | 0.05164 | 0.038341 | 0.0075277 | 0.02582 | 0.13412 | 0.022583
| Rel.Std.Dev. |  | 1.791% | 2.314% | 3.158% | 1.8% | 9.682% | 7.905% | 1.481% | 5.533% | 4.398% | 4.385%
| PDM3 |  | 14.49% | 3.899% | -5.365% | -5.127% | 5.885% | -3.711% | 0.9215% | -7.351% | 505.4% | 2.245%
| Z-Score (Absolute) |  | 1.51 | 0.33 | 0.7 | 0.67 | NA | 0.52 | 0 | 0.92 | 56.25 | 0.15
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 240b (Certified Value 0.15 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.18 | < 0.2 | 0.18003 | < 0.1 | 0.2 | < 0.2 | 14 | < 0.1 | 0.15 | 0.13 | 0.17 | 0.2 | < 5 | 0.1 | < 0.5 | 0.11 | < 1 | 0.10763 | 0.16 | 0.1972 | 0.05 | 0.15 | 0.19 | 0.16 | [5 empty] | 1
| 1 | 2 | 0.15 | < 0.2 | 0.19276 | < 0.1 | 0.1 | < 0.2 | 14 | < 0.1 | 0.15 | 0.15 | 0.15 | 0.21 | < 5 | 0.1 | < 0.5 | 0.13 | < 1 | 0.11515 | 0.15 | 0.1849 | 0.09 | 0.16 | 0.17 | 0.22 | [5 empty] | 23
| 1 | 3 | 0.18 | < 0.2 | 0.18462 | < 0.1 | 0.2 | < 0.2 | 14 | < 0.1 | 0.15 | 0.14 | 0.16 | 0.17 | < 5 | 0.1 | < 0.5 | 0.1 | < 1 | 0.14582 | 0.17 | 0.0945 | 0.13 | 0.17 | 0.16 | 0.09 | [5 empty] | 16
| 1 | 4 | 0.19 | < 0.2 | 0.19937 | 0.2 | 0.3 | < 0.2 | 14 | < 0.1 | 0.16 | 0.12 | 0.12 | 0.18 | < 5 | 0.1 | < 0.5 | 0.1 | < 1 | 0.13899 | 0.17 | 0.1602 | 0.08 | 0.14 | 0.18 | 0.15 | [5 empty] | 0.15299
| 1 | 5 | 0.16 | < 0.2 | 0.19528 | 0.1 | 0.2 | < 0.2 | 14 | 0.1 | 0.14 | 0.12 | 0.14 | 0.17 | < 5 | 0.1 | < 0.5 | 0.11 | < 1 | 0.12376 | 0.15 | 0.1602 | 0.1 | 0.15 | 0.17 | 0.14 | [5 empty] | 62
| 1 | 6 | 0.15 | < 0.2 | 0.20638 | 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.16 | 0.15 | 0.14 | 0.15 | < 5 | 0.1 | < 0.5 | 0.13 | < 1 | 0.12897 | 0.15 | 0.115 | 0.11 | 0.17 | 0.17 | 0.16
| Mean |  | 0.16833 |  | 0.19307 | 0.13333 | 0.2 |  | 14.1667 | 0.1 | 0.15167 | 0.135 | 0.14667 | 0.18 |  | 0.1 |  | 0.11333 |  | 0.12672 | 0.15833 | 0.152 | 0.093333 | 0.15667 | 0.17333 | 0.15333
| Median |  | 0.17 |  | 0.19402 | 0.1 | 0.2 |  | 14 | 0.1 | 0.15 | 0.135 | 0.145 | 0.175 |  | 0.1 |  | 0.11 |  | 0.12637 | 0.155 | 0.1602 | 0.095 | 0.155 | 0.17 | 0.155
| Std Dev. |  | 0.017224 |  | 0.0096264 | 0.057735 | 0.070711 |  | 0.40825 |  | 0.0075277 | 0.013784 | 0.017512 | 0.021909 |  | 1.5202e-17 |  | 0.013663 |  | 0.014335 | 0.0098319 | 0.039838 | 0.027325 | 0.012111 | 0.010328 | 0.041793
| Rel.Std.Dev. |  | 10.23% |  | 4.986% | 43.3% | 35.36% |  | 2.882% |  | 4.963% | 10.21% | 11.94% | 12.17% |  | 1.52e-14% |  | 12.06% |  | 11.31% | 6.21% | 26.21% | 29.28% | 7.73% | 5.958% | 27.26%
| PDM3 |  | 10.03% |  | 26.2% | -12.85% | 30.73% |  | 9160% | -34.63% | -0.8619% | -11.76% | -4.13% | 17.66% |  | -34.63% |  | -25.92% |  | -17.17% | 3.496% | -0.644% | -38.99% | 2.406% | 13.3% | 0.2275%
| Z-Score (Absolute) |  | 0.4 | 1.33 | 1.02 | 1.54 | NA | 1.33 | NA | NA | 0.03 | 0.45 | 0.15 | 0.69 | 59.42 | NA | 2.46 | 1 | 8.79 | 0.66 | 0.14 | 0.02 | 1.5 | 0.1 | 0.52 | 0.02
Results from laboratories 1.10, 1.13 and 1.22 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 240b (Certified Value 1.45 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.9 | 1.44 | 1.34 | 1.1364 | 1.3863 | 1.4 | 1.41 | 1.54 | 1.38 | 1.4 | 1.35 | 1.45 | 1.36 | 1.33 | 1.19 | 1.2 | 1.46 | 1.5 | 1.3903 | 1.47 | 1.7159 | 1.8 | 1.5 | 1.5 | 1.6 |  |  |  |  | 1
| 1 | 2 | 2.7 | 1.42 | 1.34 | 1.1756 | 1.3974 | 1.5 | 1.62 | 1.57 | 1.36 | 1.4 | 1.34 | 1.5 | 1.42 | 1.36 | 1.41 | 1.3 | 1.45 | 1.5 | 1.4291 | 1.45 | 1.7503 | 1.7 | 1.4 | 1.4 | 1.5 |  |  |  |  | 24
| 1 | 3 | 2.9 | 1.4 | 1.33 | 1.147 | 1.4305 | 1.5 | 1.47 | 1.53 | 1.42 | 1.4 | 1.29 | 1.51 | 1.36 | 1.34 | 1.32 | 1.2 | 1.52 | 1.4 | 1.5143 | 1.46 | 1.7418 | 1.7 | 1.5 | 1.5 | 1.5 |  |  |  |  | 16
| 1 | 4 | 2.5 | 1.41 | 1.26 | 1.1712 | 1.4524 | 1.5 | 1.53 | 1.6 | 1.38 | 1.4 | 1.27 | 1.42 | 1.39 | 1.42 | 1.36 | 1.2 | 1.57 | 1.5 | 1.4722 | 1.46 | 1.7677 | 1.7 | 1.4 | 1.5 | 1.6 |  |  |  |  | 1.4503
| 1 | 5 | 1.3 | 1.36 | 1.27 | 1.1506 | 1.4462 | 1.6 | 1.74 | 1.64 | 1.34 | 1.5 | 1.33 | 1.45 | 1.36 | 1.34 | 1.24 | 1.2 | 1.48 | 1.4 | 1.4512 | 1.43 | 1.694 | 1.6 | 1.4 | 1.3 | 1.4 |  |  |  |  | 63
| 1 | 6 | 1.8 | 1.46 | 1.37 | 1.1511 | 1.4357 | 1.5 | 1.76 | 1.61 | 1.36 | 1.4 | 1.35 | 1.27 | 1.37 | 1.31 | 1.4 | 1.3 | 1.54 | 1.4 | 1.5443 | 1.44 | 1.6622 | 1.7 | 1.5 | 1.4 | 1.5
| Mean |  | 2.35 | 1.415 | 1.3183 | 1.1553 | 1.4248 | 1.5 | 1.5883 | 1.5817 | 1.3733 | 1.4167 | 1.3217 | 1.4333 | 1.3767 | 1.35 | 1.32 | 1.2333 | 1.5033 | 1.45 | 1.4669 | 1.4517 | 1.722 | 1.7 | 1.45 | 1.4333 | 1.5167
| Median |  | 2.6 | 1.415 | 1.335 | 1.1508 | 1.4331 | 1.5 | 1.575 | 1.585 | 1.37 | 1.4 | 1.335 | 1.45 | 1.365 | 1.34 | 1.34 | 1.2 | 1.5 | 1.45 | 1.4617 | 1.455 | 1.7288 | 1.7 | 1.45 | 1.45 | 1.5
| Std Dev. |  | 0.65651 | 0.034496 | 0.043551 | 0.015029 | 0.026841 | 0.063246 | 0.14331 | 0.042622 | 0.027325 | 0.040825 | 0.033714 | 0.086872 | 0.024221 | 0.037947 | 0.088769 | 0.05164 | 0.04761 | 0.054772 | 0.056221 | 0.01472 | 0.039188 | 0.063246 | 0.054772 | 0.08165 | 0.075277
| Rel.Std.Dev. |  | 27.94% | 2.438% | 3.303% | 1.301% | 1.884% | 4.216% | 9.022% | 2.695% | 1.99% | 2.882% | 2.551% | 6.061% | 1.759% | 2.811% | 6.725% | 4.187% | 3.167% | 3.777% | 3.833% | 1.014% | 2.276% | 3.72% | 3.777% | 5.696% | 4.963%
| PDM3 |  | 62.03% | -2.434% | -9.1% | -20.34% | -1.761% | 3.427% | 9.517% | 9.058% | -5.307% | -2.319% | -8.87% | -1.17% | -5.077% | -6.916% | -8.985% | -14.96% | 3.656% | -0.021% | 1.146% | 0.09391% | 18.73% | 17.22% | -0.021% | -1.17% | 4.576%
| Z-Score (Absolute) |  | 8.83 | 0.18 | 1.11 | 2.68 | 0.08 | 0.64 | 1.49 | 1.43 | 0.58 | 0.16 | 1.08 | 0 | 0.55 | 0.8 | 1.09 | 1.93 | 0.67 | 0.16 | 0.32 | 0.18 | 2.78 | 2.57 | 0.16 | 0 | 0.8
| Analytical results for Ti in OREAS 240b (Certified Value 0.543 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.54 | 0.5362 | 0.54907 | 0.56519 | 0.39 | 0.5785 | 0.549 | 0.5364 | 0.51 | 0.557 | 0.546 | 0.533 | 0.547 | 0.543 | 0.53 | 0.5607 | 0.539 | 0.546 | 0.61 | 0.52055 | 0.534 | 0.56776 | 0.541 | 0.46569 | 0.53 | 0.56 | 0.55 | 0.59 |  | 1
| 1 | 2 | 0.56 | 0.5475 | 0.54205 | 0.56914 | 0.254 | 0.5935 | 0.549 | 0.5338 | 0.52 | 0.554 | 0.565 | 0.555 | 0.559 | 0.535 | 0.51 | 0.5623 | 0.537 | 0.541 | 0.57 | 0.51451 | 0.536 | 0.5636 | 0.552 | 0.47679 | 0.54 | 0.54 | 0.54 | 0.6 |  | 25
| 1 | 3 | 0.56 | 0.5465 | 0.54624 | 0.56905 | 0.362 | 0.5863 | 0.55 | 0.5518 | 0.51 | 0.551 | 0.552 | 0.553 | 0.539 | 0.537 | 0.52 | 0.5675 | 0.542 | 0.519 | 0.586 | 0.53072 | 0.531 | 0.5398 | 0.55 | 0.46979 | 0.53 | 0.55 | 0.55 | 0.59 |  | 16
| 1 | 4 | 0.55 | 0.552 | 0.54829 | 0.56237 | 0.515 | 0.5905 | 0.547 | 0.5446 | 0.51 | 0.555 | 0.548 | 0.555 | 0.544 | 0.544 | 0.53 | 0.5669 | 0.529 | 0.553 | 0.52 | 0.53455 | 0.531 | 0.55829 | 0.554 | 0.47419 | 0.54 | 0.55 | 0.55 | 0.57 |  | 0.54344
| 1 | 5 | 0.53 | 0.5247 | 0.54313 | 0.55922 | 0.597 | 0.5917 | 0.537 | 0.522 | 0.52 | 0.562 | 0.571 | 0.553 | 0.52 | 0.547 | 0.51 | 0.5612 | 0.524 | 0.55 | 0.649 | 0.53954 | 0.534 | 0.55343 | 0.55 | 0.46199 | 0.54 | 0.54 | 0.54 | 0.61 |  | 64
| 1 | 6 | 0.55 | 0.5532 | 0.54375 | 0.5593 | 0.612 | 0.5845 | 0.533 | 0.5222 | 0.52 | 0.568 | 0.568 | 0.538 | 0.526 | 0.539 | 0.51 | 0.565 | 0.525 | 0.546 | 0.589 | 0.53707 | 0.536 | 0.55405 | 0.554 | 0.47029 | 0.53 | 0.53 | 0.54 | 0.59
| Mean |  | 0.54833 | 0.54335 | 0.54542 | 0.56405 | 0.455 | 0.5875 | 0.54417 | 0.53513 | 0.515 | 0.55783 | 0.55833 | 0.54783 | 0.53917 | 0.54083 | 0.51833 | 0.56393 | 0.53267 | 0.5425 | 0.58733 | 0.52949 | 0.53367 | 0.55616 | 0.55017 | 0.46979 | 0.535 | 0.545 | 0.545 | 0.59167
| Median |  | 0.55 | 0.547 | 0.54499 | 0.56378 | 0.4525 | 0.5884 | 0.548 | 0.5351 | 0.515 | 0.556 | 0.5585 | 0.553 | 0.5415 | 0.541 | 0.515 | 0.56365 | 0.533 | 0.546 | 0.5875 | 0.53263 | 0.534 | 0.55617 | 0.551 | 0.47004 | 0.535 | 0.545 | 0.545 | 0.59
| Std Dev. |  | 0.01169 | 0.010936 | 0.0028834 | 0.0044907 | 0.14261 | 0.0055512 | 0.0072778 | 0.011921 | 0.0054772 | 0.0061779 | 0.01093 | 0.0097245 | 0.014275 | 0.0045789 | 0.0098319 | 0.0029412 | 0.0076594 | 0.012211 | 0.042791 | 0.009898 | 0.0022509 | 0.0097367 | 0.0048339 | 0.0054086 | 0.0054772 | 0.010488 | 0.0054772 | 0.013292
| Rel.Std.Dev. |  | 2.132% | 2.013% | 0.5287% | 0.7962% | 31.34% | 0.9449% | 1.337% | 2.228% | 1.064% | 1.107% | 1.958% | 1.775% | 2.648% | 0.8466% | 1.897% | 0.5216% | 1.438% | 2.251% | 7.286% | 1.869% | 0.4218% | 1.751% | 0.8786% | 1.151% | 1.024% | 1.924% | 1.005% | 2.246%
| PDM3 |  | 0.9012% | -0.01584% | 0.3652% | 3.792% | -16.27% | 8.108% | 0.1344% | -1.528% | -5.233% | 2.649% | 2.741% | 0.8092% | -0.7856% | -0.4789% | -4.619% | 3.772% | -1.982% | -0.1723% | 8.078% | -2.566% | -1.798% | 2.341% | 1.239% | -13.55% | -1.552% | 0.2878% | 0.2878% | 8.875%
| Z-Score (Absolute) |  | 0.22 | 0.07 | 0.05 | 1.17 | 5.37 | 2.57 | 0.02 | 0.57 | 1.77 | 0.79 | 0.82 | 0.19 | 0.32 | 0.22 | 1.57 | 1.16 | 0.71 | 0.12 | 2.56 | 0.91 | 0.65 | 0.69 | 0.33 | 4.49 | 0.57 | 0.02 | 0.02 | 2.82
| Analytical results for Tl in OREAS 240b (Certified Value 0.34 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.35 | 0.35 | 0.35379 | 0.33 | 0.34 | 0.35 | 0.2 | 0.31 | 0.37 | 0.32 | 0.33 | < 5 | 0.34 | < 0.5 | 0.34 | 0.59 | 0.3 | 0.34739 | 0.34 | 0.3661 | 0.31 | 0.34 | 0.33 | 0.38 | [5 empty] | 1
| 1 | 2 | 0.34 | 0.36 | 0.35618 | 0.34 | 0.36 | 0.35 | 0.1 | 0.32 | 0.38 | 0.35 | 0.35 | < 5 | 0.35 | < 0.5 | 0.34 | 1.65 | 0.35 | 0.34369 | 0.34 | 0.3622 | 0.3 | 0.33 | 0.32 | 0.36 | [5 empty] | 26
| 1 | 3 | 0.35 | 0.35 | 0.35805 | 0.34 | 0.36 | 0.37 | 0.2 | 0.3 | 0.37 | 0.32 | 0.33 | < 5 | 0.33 | < 0.5 | 0.35 | 0.51 | 0.3 | 0.37331 | 0.34 | 0.359 | 0.31 | 0.35 | 0.33 | 0.36 | [5 empty] | 16
| 1 | 4 | 0.35 | 0.36 | 0.36267 | 0.33 | 0.39 | 0.35 | 0.2 | 0.32 | 0.34 | 0.33 | 0.37 | < 5 | 0.34 | < 0.5 | 0.35 | 0.53 | 0.3 | 0.35581 | 0.36 | 0.365 | 0.3 | 0.33 | 0.34 | 0.38 | [5 empty] | 0.34229
| 1 | 5 | 0.34 | 0.32 | 0.35768 | 0.36 | 0.37 | 0.36 | 0.3 | 0.31 | 0.37 | 0.31 | 0.33 | < 5 | 0.33 | < 0.5 | 0.35 | 0.74 | 0.3 | 0.35263 | 0.36 | 0.3493 | 0.3 | 0.33 | 0.31 | 0.34 | [5 empty] | 65
| 1 | 6 | 0.35 | 0.35 | 0.35142 | 0.33 | 0.37 | 0.36 | 0.4 | 0.33 | 0.31 | 0.34 | 0.33 | < 5 | 0.36 | < 0.5 | 0.35 | 0.87 | 0.3 | 0.37824 | 0.38 | 0.3578 | 0.3 | 0.33 | 0.31 | 0.36
| Mean |  | 0.34667 | 0.34833 | 0.35663 | 0.33833 | 0.365 | 0.35667 | 0.23333 | 0.315 | 0.35667 | 0.32833 | 0.34 |  | 0.34167 |  | 0.34667 | 0.815 | 0.30833 | 0.35851 | 0.35333 | 0.3599 | 0.30333 | 0.335 | 0.32333 | 0.36333
| Median |  | 0.35 | 0.35 | 0.35693 | 0.335 | 0.365 | 0.355 | 0.2 | 0.315 | 0.37 | 0.325 | 0.33 |  | 0.34 |  | 0.35 | 0.665 | 0.3 | 0.35422 | 0.35 | 0.3606 | 0.3 | 0.33 | 0.325 | 0.36
| Std Dev. |  | 0.005164 | 0.01472 | 0.0038734 | 0.01169 | 0.016432 | 0.008165 | 0.10328 | 0.010488 | 0.026583 | 0.01472 | 0.016733 |  | 0.01169 |  | 0.005164 | 0.43145 | 0.020412 | 0.014096 | 0.01633 | 0.0061208 | 0.005164 | 0.0083666 | 0.012111 | 0.015055
| Rel.Std.Dev. |  | 1.49% | 4.226% | 1.086% | 3.455% | 4.502% | 2.289% | 44.26% | 3.33% | 7.453% | 4.483% | 4.922% |  | 3.422% |  | 1.49% | 52.94% | 6.62% | 3.932% | 4.622% | 1.701% | 1.702% | 2.497% | 3.746% | 4.144%
| PDM3 |  | 1.28% | 1.767% | 4.191% | -1.155% | 6.636% | 4.202% | -31.83% | -7.972% | 4.202% | -4.076% | -0.6677% |  | -0.1808% |  | 1.28% | 138.1% | -9.919% | 4.741% | 3.228% | 5.146% | -11.38% | -2.128% | -5.537% | 6.149%
| Z-Score (Absolute) |  | 0 | 0.09 | 0.57 | 0.47 | 1.04 | 0.57 | NA | 1.8 | 0.57 | 1.04 | 0.38 | 122.58 | 0.28 | 5.5 | 0 | 26.66 | 2.18 | 0.67 | 0.38 | 0.75 | 2.47 | 0.66 | 1.33 | 0.95
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 240b (Certified Value 0.29 ppm) | [30 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | [22 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [22 empty] | 2
3
| 1 | 1 | 0.3 | 0.26 | 0.26671 | 0.3 | 0.29 | 0.3 | 0.31149 | [22 empty] | 1
| 1 | 2 | 0.3 | 0.28 | 0.2712 | 0.3 | 0.27 | 0.3 | 0.30431 | [22 empty] | 27
| 1 | 3 | 0.31 | 0.28 | 0.27379 | 0.3 | 0.29 | 0.3 | 0.32574 | [22 empty] | 16
| 1 | 4 | 0.3 | 0.26 | 0.28306 | 0.3 | 0.29 | 0.25 | 0.31662 | [22 empty] | 0.29292
| 1 | 5 | 0.31 | 0.26 | 0.27938 | 0.3 | 0.28 | 0.3 | 0.32136 | [22 empty] | 66
| 1 | 6 | 0.32 | 0.29 | 0.27676 | 0.3 | 0.29 | 0.3 | 0.33473
| Mean |  | 0.30667 | 0.27167 | 0.27515 | 0.3 | 0.285 | 0.29167 | 0.31904
| Median |  | 0.305 | 0.27 | 0.27528 | 0.3 | 0.29 | 0.3 | 0.31899
| Std Dev. |  | 0.008165 | 0.013292 | 0.00586 | 0 | 0.0083666 | 0.020412 | 0.010736
| Rel.Std.Dev. |  | 2.662% | 4.893% | 2.13% | 0% | 2.936% | 6.999% | 3.365%
| PDM3 |  | 4.693% | -7.256% | -6.067% | 2.417% | -2.704% | -0.4281% | 8.917%
| Z-Score (Absolute) |  | 0.83 | 0.75 | 0.6 | NA | 0.15 | 0.15 | 1.39
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 240b (Certified Value 0.42 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.6 | 0.44 | 0.42 | 0.3286 | 0.42849 | 0.4 | 0.46 | 0.42 | 0.44 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.35 | 0.4 | 0.43 | 0.5 | 0.3982 | 0.4 | 0.4517 | 0.44 | 0.5 | 0.4 | 0.476 |  |  |  | 1
| 1 | 2 | 0.6 | 0.43 | 0.42 | 0.32225 | 0.42332 | 0.4 | 0.53 | 0.44 | 0.44 | 0.4 | 0.4 | 0.5 | 0.4 | 0.4 | < 10 | 0.41 | 0.4 | 0.43 | 0.5 | 0.40469 | 0.4 | 0.4528 | 0.424 | 0.5 | 0.4 | 0.453 |  |  |  | 28
| 1 | 3 | 0.6 | 0.44 | 0.42 | 0.33739 | 0.43629 | 0.4 | 0.51 | 0.41 | 0.43 | 0.4 | 0.4 | 0.5 | 0.4 | 0.4 | < 10 | 0.38 | 0.4 | 0.45 | 0.5 | 0.43102 | 0.4 | 0.4602 | 0.435 | 0.5 | 0.4 | 0.432 |  |  |  | 16
| 1 | 4 | 0.6 | 0.43 | 0.39 | 0.32624 | 0.44295 | 0.4 | 0.5 | 0.44 | 0.43 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.39 | 0.4 | 0.41 | 0.5 | 0.42204 | 0.4 | 0.4869 | 0.429 | 0.4 | 0.4 | 0.473 |  |  |  | 0.42353
| 1 | 5 | 0.5 | 0.42 | 0.4 | 0.31805 | 0.47895 | 0.4 | 0.52 | 0.44 | 0.42 | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.35 | 0.4 | 0.45 | 0.4 | 0.41671 | 0.4 | 0.4487 | 0.424 | 0.5 | 0.4 | 0.418 |  |  |  | 67
| 1 | 6 | 0.5 | 0.45 | 0.42 | 0.32709 | 0.50958 | 0.4 | 0.56 | 0.42 | 0.42 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.4 | 0.4 | 0.46 | 0.4 | 0.42965 | 0.4 | 0.4386 | 0.434 | 0.5 | 0.4 | 0.476
| Mean |  | 0.56667 | 0.435 | 0.41167 | 0.3266 | 0.45326 | 0.4 | 0.51333 | 0.42833 | 0.43 | 0.41667 | 0.4 | 0.43333 | 0.4 | 0.4 |  | 0.38 | 0.4 | 0.43833 | 0.46667 | 0.41705 | 0.4 | 0.45648 | 0.431 | 0.48333 | 0.4 | 0.45467
| Median |  | 0.6 | 0.435 | 0.42 | 0.32667 | 0.43962 | 0.4 | 0.515 | 0.43 | 0.43 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 |  | 0.385 | 0.4 | 0.44 | 0.5 | 0.41938 | 0.4 | 0.45225 | 0.4315 | 0.5 | 0.4 | 0.463
| Std Dev. |  | 0.05164 | 0.010488 | 0.013292 | 0.0065213 | 0.033871 | 6.0809e-17 | 0.033267 | 0.013292 | 0.0089443 | 0.040825 | 6.0809e-17 | 0.05164 | 6.0809e-17 | 6.0809e-17 |  | 0.025298 | 6.0809e-17 | 0.018348 | 0.05164 | 0.013319 | 6.0809e-17 | 0.016467 | 0.0064498 | 0.040825 | 6.0809e-17 | 0.024929
| Rel.Std.Dev. |  | 9.113% | 2.411% | 3.229% | 1.997% | 7.473% | 1.52e-14% | 6.481% | 3.103% | 2.08% | 9.798% | 1.52e-14% | 11.92% | 1.52e-14% | 1.52e-14% |  | 6.657% | 1.52e-14% | 4.186% | 11.07% | 3.194% | 1.52e-14% | 3.607% | 1.496% | 8.447% | 1.52e-14% | 5.483%
| PDM3 |  | 33.8% | 2.709% | -2.8% | -22.88% | 7.021% | -5.555% | 21.2% | 1.135% | 1.528% | -1.62% | -5.555% | 2.315% | -5.555% | -5.555% |  | -10.28% | -5.555% | 3.496% | 10.19% | -1.529% | -5.555% | 7.781% | 1.764% | 14.12% | -5.555% | 7.352%
| Z-Score (Absolute) |  | 3.18 | 0.13 | 0.4 | 2.37 | 0.56 | 0.67 | 1.95 | 0.02 | 0.02 | 0.29 | 0.67 | 0.1 | 0.67 | 0.67 | 105.67 | 1.14 | 0.67 | 0.21 | 0.87 | 0.28 | 0.67 | 0.63 | 0.04 | 1.25 | 0.67 | 0.59
| Analytical results for V in OREAS 240b (Certified Value 232 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 233 | 230 | 207.31 | 221.5854 | 232.1608 | 179 | 281 | 217 | 228 | 222 | 225 | 231 | 223 | 232 | 230 | 245 | 231 | 232 | 239 | 194.28 | 228.5464 | 230 | 240.6788 | 231 | 216.4535 | 239 | 233 | 245 | 252 | 1
| 1 | 2 | 235 | 240 | 211.81 | 224.3803 | 234.734 | 104 | 282 | 214 | 232 | 222 | 225 | 234 | 235 | 234 | 227 | 236 | 232 | 232 | 241 | 176.46 | 225.4276 | 225 | 239.0539 | 234 | 218.6591 | 236 | 225 | 244 | 238 | 29
| 1 | 3 | 229 | 225 | 208.26 | 222.4115 | 234.7113 | 146 | 284 | 225 | 234 | 219 | 227 | 230 | 234 | 231 | 225 | 239 | 231 | 236 | 232 | 197.9 | 233.0268 | 225 | 235.7971 | 227 | 217.8511 | 234 | 233 | 245 | 213 | 16
| 1 | 4 | 229 | 232 | 209.47 | 223.9582 | 230.7493 | 210 | 286 | 230 | 199 | 220 | 229 | 229 | 233 | 234 | 230 | 239 | 233 | 230 | 226 | 178.2 | 236.8172 | 225 | 239.5583 | 236 | 216.8059 | 236 | 228 | 244 | 237 | 231.5352
| 1 | 5 | 240 | 223 | 207.9 | 224.2663 | 229.1735 | 251 | 284 | 225 | 223 | 221 | 239 | 236 | 232 | 224 | 227 | 236 | 233 | 223 | 243 | 206.85 | 236.8356 | 230 | 239.9844 | 234 | 215.2504 | 231 | 221 | 244 | 220 | 68
| 1 | 6 | 246 | 243 | 208.53 | 221.2623 | 230.5221 | 222 | 286 | 231 | 226 | 225 | 242 | 237 | 226 | 226 | 227 | 231 | 232 | 225 | 242 | 191.88 | 238.6204 | 230 | 242.2027 | 233 | 212.5537 | 234 | 223 | 244 | 235
| Mean |  | 235.3333 | 232.1667 | 208.88 | 222.9773 | 232.0085 | 185.3333 | 283.8333 | 223.6667 | 223.6667 | 221.5 | 231.1667 | 232.8333 | 230.5 | 230.1667 | 227.6667 | 237.6667 | 232 | 229.6667 | 237.1667 | 190.9283 | 233.2123 | 227.5 | 239.5459 | 232.5 | 216.2623 | 235 | 227.1667 | 244.3333 | 232.5
| Median |  | 234 | 231 | 208.395 | 223.1848 | 231.455 | 194.5 | 284 | 225 | 227 | 221.5 | 228 | 232.5 | 232.5 | 231.5 | 227 | 237.5 | 232 | 231 | 240 | 193.08 | 234.922 | 227.5 | 239.7713 | 233.5 | 216.6297 | 235 | 226.5 | 244 | 236
| Std Dev. |  | 6.6533 | 7.9854 | 1.6041 | 1.3994 | 2.306 | 53.791 | 2.0412 | 6.8605 | 12.7227 | 2.0736 | 7.4409 | 3.3116 | 4.8477 | 4.2151 | 1.9664 | 4.6332 | 0.89443 | 4.8442 | 6.7355 | 11.7097 | 5.2493 | 2.7386 | 2.1368 | 3.1464 | 2.1623 | 2.6833 | 5.0761 | 0.5164 | 13.9535
| Rel.Std.Dev. |  | 2.827% | 3.44% | 0.7679% | 0.6276% | 0.9939% | 29.02% | 0.7192% | 3.067% | 5.688% | 0.9362% | 3.219% | 1.422% | 2.103% | 1.831% | 0.8637% | 1.949% | 0.3855% | 2.109% | 2.84% | 6.133% | 2.251% | 1.204% | 0.892% | 1.353% | 0.9999% | 1.142% | 2.235% | 0.2113% | 6.002%
| PDM3 |  | 1.64% | 0.2727% | -9.785% | -3.696% | 0.2044% | -19.95% | 22.59% | -3.398% | -3.398% | -4.334% | -0.1592% | 0.5607% | -0.4471% | -0.5911% | -1.671% | 2.648% | 0.2008% | -0.807% | 2.432% | -17.54% | 0.7244% | -1.743% | 3.46% | 0.4167% | -6.596% | 1.496% | -1.887% | 5.528% | 0.4167%
| Z-Score (Absolute) |  | 0.7 | 0.17 | 3.76 | 1.38 | 0.14 | 7.73 | 8.88 | 1.26 | 1.26 | 1.63 | 0 | 0.28 | 0.11 | 0.17 | 0.59 | 1.1 | 0.14 | 0.25 | 1.01 | 6.78 | 0.34 | 0.62 | 1.41 | 0.22 | 2.51 | 0.65 | 0.67 | 2.22 | 0.22
| Analytical results for W in OREAS 240b (Certified Value 47.1 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 52.7 | 45.4 | 47.87 | 36.2642 | 46.5022 | 5.9 | 44 | 44.7 | 45.1 | 47 | 44.2 | 49.3 | 48.1 | 42.5 | 46.5 | 43 | 39.8 | 42.5 | 47 | 54 | 44.8647 | 51.1 | 50.0678 | 50.3 | 48 | 48.8 | 53 |  |  | 1
| 1 | 2 | 52.7 | 47 | 48.87 | 36.5562 | 48.1161 | 1.4 | 47.1 | 45.2 | 46.5 | 46 | 42.6 | 50.1 | 49.2 | 44.6 | 45.7 | 42 | 43.5 | 43.9 | 46 | 55 | 44.8909 | 49.7 | 51.1254 | 50.1 | 46.8 | 48.3 | 50.2 |  |  | 30
| 1 | 3 | 53.3 | 46.9 | 48.31 | 36.6318 | 49.3227 | 2.1 | 45.2 | 44.6 | 45.6 | 45 | 45 | 48 | 48.1 | 42.7 | 47.8 | 42 | 41 | 43.2 | 46 | 51 | 48.1084 | 51.2 | 50.3547 | 50.2 | 47.3 | 49.4 | 48.4 |  |  | 16
| 1 | 4 | 52.6 | 46.5 | 49.44 | 36.7764 | 49.2323 | 28.8 | 46.1 | 44.8 | 44.9 | 46 | 42 | 47.1 | 47.5 | 44.2 | 49.3 | 42 | 41.7 | 42.7 | 53 | 52 | 46.4778 | 51.2 | 50.4184 | 49.7 | 45.8 | 50.4 | 52.7 |  |  | 47.1217
| 1 | 5 | 52.3 | 45.3 | 48.12 | 36.7786 | 48.9029 | 44.6 | 46.2 | 44.6 | 43.5 | 47 | 43.9 | 48.9 | 49 | 42.3 | 46.4 | 41 | 38.7 | 42.2 | 47 | 52 | 46.6286 | 51.2 | 48.9025 | 50 | 45.1 | 48.9 | 48.2 |  |  | 69
| 1 | 6 | 52.3 | 47.1 | 49.82 | 36.3512 | 48.7788 | 39 | 52.5 | 44.4 | 44.5 | 46 | 46.1 | 50.4 | 48.2 | 42.4 | 45 | 41 | 43.4 | 42.8 | 46 | 50 | 48.0099 | 52.3 | 48.2368 | 50.8 | 46.3 | 49.5 | 49.9
| Mean |  | 52.65 | 46.3667 | 48.7383 | 36.5597 | 48.4758 | 20.3 | 46.85 | 44.7167 | 45.0167 | 46.1667 | 43.9667 | 48.9667 | 48.35 | 43.1167 | 46.7833 | 41.8333 | 41.35 | 42.8833 | 47.5 | 52.3333 | 46.4967 | 51.1167 | 49.8509 | 50.1833 | 46.55 | 49.2167 | 50.4
| Median |  | 52.65 | 46.7 | 48.59 | 36.594 | 48.8409 | 17.35 | 46.15 | 44.65 | 45 | 46 | 44.05 | 49.1 | 48.15 | 42.6 | 46.45 | 42 | 41.35 | 42.75 | 46.5 | 52 | 46.5532 | 51.2 | 50.2112 | 50.15 | 46.55 | 49.15 | 50.05
| Std Dev. |  | 0.36742 | 0.81404 | 0.77453 | 0.21487 | 1.0572 | 19.5358 | 2.9602 | 0.27142 | 1.0128 | 0.75277 | 1.5108 | 1.2549 | 0.63482 | 1.0108 | 1.5459 | 0.75277 | 1.9233 | 0.59805 | 2.7386 | 1.8619 | 1.4246 | 0.82805 | 1.0726 | 0.3656 | 1.0445 | 0.72503 | 2.0582
| Rel.Std.Dev. |  | 0.6979% | 1.756% | 1.589% | 0.5877% | 2.181% | 96.24% | 6.319% | 0.607% | 2.25% | 1.631% | 3.436% | 2.563% | 1.313% | 2.344% | 3.304% | 1.799% | 4.651% | 1.395% | 5.766% | 3.558% | 3.064% | 1.62% | 2.152% | 0.7285% | 2.244% | 1.473% | 4.084%
| PDM3 |  | 11.73% | -1.602% | 3.431% | -22.41% | 2.874% | -56.92% | -0.5767% | -5.104% | -4.467% | -2.027% | -6.696% | 3.915% | 2.607% | -8.499% | -0.7181% | -11.22% | -12.25% | -8.995% | 0.8028% | 11.06% | -1.326% | 8.478% | 5.792% | 6.497% | -1.213% | 4.446% | 6.957%
| Z-Score (Absolute) |  | 1.63 | 0.12 | 0.54 | 2.83 | 0.47 | 7.34 | 0.02 | 0.57 | 0.49 | 0.17 | 0.78 | 0.61 | 0.43 | 1.02 | 0 | 1.37 | 1.51 | 1.08 | 0.2 | 1.54 | 0.08 | 1.2 | 0.85 | 0.94 | 0.06 | 0.67 | 1
| Analytical results for Y in OREAS 240b (Certified Value 18.3 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 17.3 | 18.03 | 15.99 | 16.6658 | 18.1281 | 17.5 | 16.5 | 19.1 | 17.54 | 18 | 18.8 | 18.5 | 19.4 | 18.4 | 19.4 | 18 | 16.2 | 17.2 | 18 | 19.5 | 18.9806 | 19.5 | 18.0456 | 19.3 | 19.4 | 19.1 | 19.5 |  |  | 1
| 1 | 2 | 17.7 | 18.08 | 17.03 | 16.5922 | 18.2352 | 17.2 | 18.9 | 19.1 | 17.43 | 18 | 18.9 | 19.2 | 20 | 19.6 | 18.8 | 17 | 17.5 | 17.7 | 18.2 | 19.1 | 18.9473 | 19.2 | 18.4288 | 19 | 18.1 | 18.3 | 18.6 |  |  | 31
| 1 | 3 | 17.7 | 18.34 | 16.99 | 16.914 | 18.5527 | 17.3 | 17.7 | 19.3 | 17.9 | 17 | 19.3 | 18.4 | 20 | 18.7 | 19.6 | 18 | 17 | 17.4 | 18.1 | 18.6 | 18.9032 | 19.5 | 18.3187 | 19.2 | 19.3 | 19.4 | 17.4 |  |  | 16
| 1 | 4 | 17.5 | 18.34 | 15.73 | 16.8806 | 18.2483 | 17.6 | 18.3 | 19.4 | 17.54 | 17 | 18.9 | 18.5 | 19 | 19 | 20.3 | 18 | 17.7 | 17.2 | 17.8 | 18.7 | 19.365 | 19.2 | 18.2556 | 18.8 | 18.1 | 19.6 | 20 |  |  | 18.2682
| 1 | 5 | 16.8 |  | 16 | 16.5253 | 18.0507 | 18.6 | 18.3 | 19.3 | 17.16 | 18 | 19.3 | 18.6 | 19.2 | 18.4 | 19.3 | 17 | 16.2 | 16.7 | 17.7 | 19.2 | 18.5438 | 19.5 | 17.6413 | 19.5 | 18.4 | 19.4 | 17 |  |  | 70
| 1 | 6 | 17.3 | 18.43 | 16.73 | 16.5495 | 18.0782 | 17 | 20.7 | 19.1 | 17.26 | 18 | 19.2 | 18.9 | 19 | 18.1 | 19.5 | 17 | 18.4 | 16.9 | 18.7 | 18.6 | 19.1568 | 19 | 16.9931 | 19.1 | 18.6 | 19.3 | 18.9
| Mean |  | 17.3833 | 18.244 | 16.4117 | 16.6879 | 18.2155 | 17.5333 | 18.4 | 19.2167 | 17.4717 | 17.6667 | 19.0667 | 18.6833 | 19.4333 | 18.7 | 19.4833 | 17.5 | 17.1667 | 17.1833 | 18.0833 | 18.95 | 18.9828 | 19.3167 | 17.9472 | 19.15 | 18.65 | 19.1833 | 18.5667
| Median |  | 17.4 | 18.34 | 16.365 | 16.629 | 18.1817 | 17.4 | 18.3 | 19.2 | 17.485 | 18 | 19.05 | 18.55 | 19.3 | 18.55 | 19.45 | 17.5 | 17.25 | 17.2 | 18.05 | 18.9 | 18.9639 | 19.35 | 18.1506 | 19.15 | 18.5 | 19.35 | 18.75
| Std Dev. |  | 0.33714 | 0.17729 | 0.57098 | 0.16942 | 0.18374 | 0.56451 | 1.39 | 0.13292 | 0.25942 | 0.5164 | 0.22509 | 0.30605 | 0.46332 | 0.53666 | 0.48751 | 0.54772 | 0.87331 | 0.35449 | 0.35449 | 0.37283 | 0.27447 | 0.2137 | 0.54381 | 0.2429 | 0.57533 | 0.46224 | 1.1708
| Rel.Std.Dev. |  | 1.939% | 0.9717% | 3.479% | 1.015% | 1.009% | 3.22% | 7.554% | 0.6917% | 1.485% | 2.923% | 1.181% | 1.638% | 2.384% | 2.87% | 2.502% | 3.13% | 5.087% | 2.063% | 1.96% | 1.967% | 1.446% | 1.106% | 3.03% | 1.268% | 3.085% | 2.41% | 6.306%
| PDM3 |  | -4.844% | -0.1323% | -10.16% | -8.65% | -0.2881% | -4.023% | 0.7216% | 5.192% | -4.36% | -3.293% | 4.371% | 2.273% | 6.378% | 2.364% | 6.652% | -4.205% | -6.03% | -5.938% | -1.012% | 3.732% | 3.912% | 5.739% | -1.757% | 4.827% | 2.09% | 5.01% | 1.634%
| Z-Score (Absolute) |  | 0.91 | 0.14 | 1.79 | 1.54 | 0.17 | 0.78 | 0 | 0.73 | 0.83 | 0.66 | 0.6 | 0.25 | 0.93 | 0.27 | 0.97 | 0.81 | 1.11 | 1.09 | 0.28 | 0.49 | 0.52 | 0.82 | 0.41 | 0.67 | 0.22 | 0.7 | 0.15
| Analytical results for Yb in OREAS 240b (Certified Value 1.87 ppm) | [30 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 2.2 | 1.87 | 1.61 | 1.8605 | 1.9 | 1.69 | 2 | 1.8 | 1.9839 | 2 | [19 empty] | 1
| 1 | 2 | 2.2 | 1.96 | 1.75 | 1.8359 | 1.9 | 1.88 | 2 | 1.8 | 1.9803 | 1.9 | [19 empty] | 32
| 1 | 3 | 2.2 | 1.89 | 1.76 | 1.7891 | 1.9 | 1.77 | 2 | 1.6 | 2.1396 | 1.9 | [19 empty] | 16
| 1 | 4 | 2.2 | 1.97 | 1.72 | 1.8371 | 1.9 | 1.83 | 2 | 1.7 | 2.1181 | 1.9 | [19 empty] | 1.8663
| 1 | 5 | 2.1 | 1.83 | 1.98 | 1.8065 | 2.1 | 1.96 | 2 | 1.6 | 2.0945 | 1.8 | [19 empty] | 71
| 1 | 6 | 2.2 | 1.98 | 1.73 | 1.8399 | 1.9 | 2.14 | 2 | 1.7 | 2.1419 | 1.8
| Mean |  | 2.1833 | 1.9167 | 1.7583 | 1.8282 | 1.9333 | 1.8783 | 2 | 1.7 | 2.0764 | 1.8833
| Median |  | 2.2 | 1.925 | 1.74 | 1.8365 | 1.9 | 1.855 | 2 | 1.7 | 2.1063 | 1.9
| Std Dev. |  | 0.040825 | 0.06186 | 0.12123 | 0.025774 | 0.08165 | 0.15791 | 0 | 0.089443 | 0.075022 | 0.075277
| Rel.Std.Dev. |  | 1.87% | 3.227% | 6.895% | 1.41% | 4.223% | 8.407% | 0% | 5.261% | 3.613% | 3.997%
| PDM3 |  | 16.99% | 2.7% | -5.784% | -2.042% | 3.593% | 0.6458% | 7.165% | -8.91% | 11.26% | 0.9137%
| Z-Score (Absolute) |  | 3.67 | 0.41 | 1.53 | 0.67 | 0.61 | 0.06 | NA | 2.24 | 2.36 | 0
Results from laboratory 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 240b (Certified Value 173 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 175 | 180 | 164.45 | 177.76 | 301.1871 | 168 | 166 | 178 | 172 | 162 | 169 | 178 | 182 | 183 | 181 | 177 | 177 | 168 | 168 | 72.14 | 159.503 | 170 | 173.6956 | 181 | 207.9719 | 162 | 181 | 162 | 182 | 1
| 1 | 2 | 177 | 181 | 167.77 | 175.6 | 177.2534 | 166 | 196 | 178 | 171 | 159 | 171 | 179 | 191 | 185 | 179 | 172 | 180 | 166 | 166 | 71.86 | 158.9307 | 170 | 172.1952 | 183 | 229.8739 | 162 | 171 | 161 | 176 | 33
| 1 | 3 | 180 | 181 | 164.1 | 177.57 | 177.3905 | 166 | 178 | 178 | 176 | 162 | 169 | 177 | 189 | 181 | 179 | 174 | 178 | 167 | 163 | 65.13 | 164.0232 | 170 | 170.4492 | 185 | 220.5477 | 164 | 178 | 163 | 167 | 16
| 1 | 4 | 177 | 184 | 167.46 | 174.15 | 175.9031 | 166 | 183 | 178 | 174 | 160 | 167 | 176 | 188 | 182 | 182 | 172 | 180 | 166 | 165 | 79.37 | 170.1633 | 170 | 176.1808 | 185 | 218.3696 | 159 | 173 | 161 | 188 | 173.4693
| 1 | 5 | 173 | 175 | 164.14 | 171.64 | 174.3126 | 164 | 181 | 175 | 166 | 163 | 180 | 182 | 190 | 176 | 182 | 170 | 176 | 166 | 172 | 80.73 | 168.9314 | 170 | 171.4213 | 184 | 212.4095 | 161 | 174 | 162 | 165 | 72
| 1 | 6 | 178 | 184 | 165.97 | 178.28 | 174.6095 | 164 | 207 | 175 | 169 | 163 | 175 | 184 | 182 | 177 | 182 | 166 | 179 | 168 | 168 | 79.26 | 168.0099 | 170 | 177.47 | 186 | 202.5694 | 163 | 180 | 160 | 183
| Mean |  | 176.6667 | 180.8333 | 165.6483 | 175.8333 | 196.776 | 165.6667 | 185.1667 | 177 | 171.3333 | 161.5 | 171.8333 | 179.3333 | 187 | 180.6667 | 180.8333 | 171.8333 | 178.3333 | 166.8333 | 167 | 74.7483 | 164.9269 | 170 | 173.5687 | 184 | 215.2903 | 161.8333 | 176.1667 | 161.5 | 176.8333
| Median |  | 177 | 181 | 165.21 | 176.585 | 176.5782 | 166 | 182 | 178 | 171.5 | 162 | 170 | 178.5 | 188.5 | 181.5 | 181.5 | 172 | 178.5 | 166.5 | 167 | 75.7 | 166.0165 | 170 | 172.9454 | 184.5 | 215.3895 | 162 | 176 | 161.5 | 179
| Std Dev. |  | 2.4221 | 3.3116 | 1.6731 | 2.5763 | 51.1669 | 1.5055 | 14.3863 | 1.5492 | 3.559 | 1.6432 | 4.8339 | 3.0768 | 4 | 3.5024 | 1.472 | 3.7103 | 1.633 | 0.98319 | 3.0984 | 6.0853 | 4.8816 | 0 | 2.7673 | 1.7889 | 9.7312 | 1.7224 | 4.0702 | 1.0488 | 9.2394
| Rel.Std.Dev. |  | 1.371% | 1.831% | 1.01% | 1.465% | 26% | 0.9088% | 7.769% | 0.8753% | 2.077% | 1.017% | 2.813% | 1.716% | 2.139% | 1.939% | 0.814% | 2.159% | 0.9157% | 0.5893% | 1.855% | 8.141% | 2.96% | 0% | 1.594% | 0.9722% | 4.52% | 1.064% | 2.31% | 0.6494% | 5.225%
| PDM3 |  | 1.843% | 4.245% | -4.509% | 1.363% | 13.44% | -4.498% | 6.743% | 2.035% | -1.231% | -6.9% | -0.9431% | 3.38% | 7.8% | 4.149% | 4.245% | -0.9431% | 2.804% | -3.825% | -3.729% | -56.91% | -4.924% | -2% | 0.05729% | 6.071% | 24.11% | -6.708% | 1.555% | -6.9% | 1.939%
| Z-Score (Absolute) |  | 0.1 | 0.58 | 1.18 | 0 | 2.42 | 1.18 | 1.08 | 0.13 | 0.52 | 1.66 | 0.46 | 0.4 | 1.29 | 0.56 | 0.58 | 0.46 | 0.29 | 1.04 | 1.02 | 11.69 | 1.26 | 0.67 | 0.26 | 0.94 | 4.56 | 1.62 | 0.04 | 1.66 | 0.12
| Analytical results for Zr in OREAS 240b (Certified Value 53 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 62.3 | 53.7 | 69.88 | 60.4795 | 56.9706 | 44 | 60 | 56.8 | 46.6 | 56 | 57.5 | 51.5 | 55.5 | 67.7 | 50.4 | 47 | 46.6 | 48.2 | 49.6 | 50.2688 | 50 | 75.9436 | 56.9 | 34.6307 | 50.6 | 86 | 50.9 | 51.1 |  | 1
| 1 | 2 | 64 | 56.1 | 71.15 | 59.4416 | 54.0949 | 17 | 61 | 56.7 | 46.1 | 56 | 57 | 54.9 | 55.6 | 57.4 | 48.5 | 45 | 51.7 | 47.6 | 49 | 50.1308 | 50 | 75.3591 | 54.6 | 35.4199 | 50.1 | 80.9 | 50.3 | 49.5 |  | 34
| 1 | 3 | 63.4 | 52.6 | 70.82 | 60.1497 | 54.1253 | 31 | 61 | 53.9 | 50.8 | 57 | 56 | 53.6 | 56.8 | 53.2 | 49.8 | 47 | 49.7 | 47.8 | 48.1 | 52.2744 | 50 | 73.7652 | 62.2 | 35.5003 | 50.5 | 84.1 | 51.3 | 46.2 |  | 16
| 1 | 4 | 62.2 | 55.7 | 67.41 | 59.4998 | 54.5323 | 50 | 62 | 53.9 | 49.3 | 55 | 56.4 | 51.8 | 53.7 | 56.3 | 55 | 46 | 48.6 | 67 | 51.9 | 52.4308 | 55 | 76.4522 | 56.7 | 34.8593 | 51 | 86.5 | 52.4 | 52.4 |  | 53.1248
| 1 | 5 | 53.3 | 49 | 67.11 | 58.6947 | 53.359 | 57 | 61 | 53.3 | 46.9 | 57 | 58.1 | 52.3 | 61.4 | 57.5 | 50.3 | 44 | 45.7 | 45 | 48.9 | 52.9644 | 55 | 73.2511 | 57.2 | 33.7959 | 52.4 | 81.2 | 51.5 | 45.8 |  | 73
| 1 | 6 | 57.7 | 53.9 | 65.39 | 58.782 | 54.1644 | 52 | 62 | 52.9 | 46.7 | 57 | 58.4 | 52.6 | 55.6 | 54.2 | 50.3 | 46 | 52.2 | 47.3 | 49.4 | 52.9828 | 55 | 75.6338 | 59.7 | 32.598 | 51.9 | 79.6 | 52.7 | 49.6
| Mean |  | 60.4833 | 53.5 | 68.6267 | 59.5079 | 54.5411 | 41.8333 | 61.1667 | 54.5833 | 47.7333 | 56.3333 | 57.2333 | 52.7833 | 56.4333 | 57.7167 | 50.7167 | 45.8333 | 49.0833 | 50.4833 | 49.4833 | 51.842 | 52.5 | 75.0675 | 57.8833 | 34.4674 | 51.0833 | 83.05 | 51.5167 | 49.1
| Median |  | 62.25 | 53.8 | 68.645 | 59.4707 | 54.1449 | 47 | 61 | 53.9 | 46.8 | 56.5 | 57.25 | 52.45 | 55.6 | 56.85 | 50.3 | 46 | 49.15 | 47.7 | 49.2 | 52.3526 | 52.5 | 75.4965 | 57.05 | 34.745 | 50.8 | 82.65 | 51.4 | 49.55
| Std Dev. |  | 4.1586 | 2.5636 | 2.324 | 0.71395 | 1.2501 | 15.0919 | 0.75277 | 1.7209 | 1.8747 | 0.8165 | 0.94375 | 1.2671 | 2.6281 | 5.1882 | 2.2158 | 1.169 | 2.6377 | 8.1698 | 1.2922 | 1.3037 | 2.7386 | 1.2717 | 2.6664 | 1.1045 | 0.88863 | 2.8863 | 0.90425 | 2.6306
| Rel.Std.Dev. |  | 6.876% | 4.792% | 3.386% | 1.2% | 2.292% | 36.08% | 1.231% | 3.153% | 3.928% | 1.449% | 1.649% | 2.401% | 4.657% | 8.989% | 4.369% | 2.551% | 5.374% | 16.18% | 2.611% | 2.515% | 5.216% | 1.694% | 4.607% | 3.205% | 1.74% | 3.475% | 1.755% | 5.358%
| PDM3 |  | 13.85% | 0.7063% | 29.18% | 12.02% | 2.666% | -21.25% | 15.14% | 2.745% | -10.15% | 6.04% | 7.734% | -0.6428% | 6.228% | 8.644% | -4.533% | -13.73% | -7.607% | -4.972% | -6.855% | -2.415% | -1.176% | 41.3% | 8.957% | -35.12% | -3.843% | 56.33% | -3.027% | -7.576%
| Z-Score (Absolute) |  | 1.22 | 0.06 | 2.58 | 1.06 | 0.23 | 1.88 | 1.34 | 0.24 | 0.9 | 0.53 | 0.68 | 0.06 | 0.55 | 0.76 | 0.4 | 1.22 | 0.68 | 0.44 | 0.61 | 0.22 | 0.11 | 3.65 | 0.79 | 3.11 | 0.34 | 4.98 | 0.27 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C952: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1026: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 240b (Certified Value 1.39 ppm) | [32 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.92 | 1.39 | 1.39 | 1.388 | 1.34 | 1.51 | 1.36 | 1.37 | 1.5 | 1.4 | 1.43 | 1.33 | 1.36 | < 10 | 1.32 | 1.2 | 1.3887 | 1.23 | 1.3 | 1.37 | 1.09 | 1.2 | 1.3753 | 1.4 | 1.421 | 1.37 | 1.41 | 1.35 | 1.43 |  |  | 1
| 1 | 2 | 1.92 | 1.37 | 1.41 | 1.4075 | 1.36 | 1.52 | 1.37 | 1.38 | 1.5 | 1.4 | 1.47 | 1.36 | 1.42 | < 10 | 1.36 | 1.2 | 1.4247 | 1.27 | 1.296 | 1.42 | 0.94 | 1.2 | 1.3811 | 1.4 | 1.443 | 1.33 | 1.41 | 1.37 | 1.4 |  |  | 18
| 1 | 3 | 1.94 | 1.36 | 1.41 | 1.4214 | 1.33 | 1.5 | 1.37 | 1.34 | 1.5 | 1.4 | 1.36 | 1.32 | 1.37 | < 10 | 1.29 | 1.3 | 1.2027 | 1.3 | 1.263 | 1.4 | 0.9 | 1.2 | 1.4255 | 1.42 | 1.489 | 1.37 | 1.47 | 1.39 | 1.45 |  |  | 16
| 1 | 4 | 1.94 | 1.41 | 1.55 | 1.3902 | 1.32 | 1.46 | 1.38 | 1.34 | 1.5 | 1.4 | 1.48 | 1.31 | 1.37 | < 10 | 1.39 | 1.3 | 1.405 | 1.28 | 1.127 | 1.48 | 0.94 | 1.2 | 1.4401 | 1.4 | 1.336 | 1.4 | 1.43 | 1.35 | 1.46 |  |  | 1.3898
| 1 | 5 | 1.88 | 1.37 | 1.47 | 1.3934 | 1.36 | 1.49 | 1.33 | 1.33 | 1.4 | 1.5 | 1.45 | 1.42 | 1.4 | < 10 | 1.4 | 1.2 | < 1 | 1.33 | 1.329 | 1.5 | 1.03 | 1.2 | 1.3777 | 1.4 | 1.481 | 1.35 | 1.44 | 1.41 | 1.42 |  |  | 75
| 1 | 6 | 1.95 | 1.39 | 1.38 | 1.4281 | 1.36 | 1.5 | 1.35 | 1.33 | 1.5 | 1.4 | 1.46 | 1.35 | 1.4 | < 10 | 1.41 | 1.3 | < 1 | 1.21 | 1.276 | 1.43 | 0.97 | 1.2 | 1.4079 | 1.4 | 1.463 | 1.33 | 1.42 | 1.37 | 1.43
| Mean |  | 1.925 | 1.3817 | 1.435 | 1.4048 | 1.345 | 1.4967 | 1.36 | 1.3483 | 1.4833 | 1.4167 | 1.4417 | 1.3483 | 1.3867 |  | 1.3617 | 1.25 | 1.3553 | 1.27 | 1.2652 | 1.4333 | 0.97833 | 1.2 | 1.4013 | 1.4033 | 1.4388 | 1.3583 | 1.43 | 1.3733 | 1.4317
| Median |  | 1.93 | 1.38 | 1.41 | 1.4005 | 1.35 | 1.5 | 1.365 | 1.34 | 1.5 | 1.4 | 1.455 | 1.34 | 1.385 |  | 1.375 | 1.25 | 1.3968 | 1.275 | 1.286 | 1.425 | 0.955 | 1.2 | 1.3945 | 1.4 | 1.453 | 1.36 | 1.425 | 1.37 | 1.43
| Std Dev. |  | 0.0251 | 0.018348 | 0.06442 | 0.017039 | 0.017607 | 0.020656 | 0.017889 | 0.02137 | 0.040825 | 0.040825 | 0.043551 | 0.039707 | 0.023381 |  | 0.047924 | 0.054772 | 0.10279 | 0.044272 | 0.071345 | 0.048854 | 0.06969 | 0 | 0.027465 | 0.008165 | 0.056194 | 0.027142 | 0.022804 | 0.023381 | 0.02137
| Rel.Std.Dev. |  | 1.304% | 1.328% | 4.489% | 1.213% | 1.309% | 1.38% | 1.315% | 1.585% | 2.752% | 2.882% | 3.021% | 2.945% | 1.686% |  | 3.519% | 4.382% | 7.584% | 3.486% | 5.639% | 3.408% | 7.123% | 0% | 1.96% | 0.5818% | 3.906% | 1.998% | 1.595% | 1.702% | 1.493%
| PDM3 |  | 38.51% | -0.5823% | 3.255% | 1.08% | -3.221% | 7.693% | -2.141% | -2.981% | 6.733% | 1.936% | 3.735% | -2.981% | -0.2225% |  | -2.021% | -10.06% | -2.482% | -8.617% | -8.965% | 3.135% | -29.6% | -13.65% | 0.8279% | 0.9768% | 3.531% | -2.261% | 2.896% | -1.182% | 3.015%
| Z-Score (Absolute) |  | 8.38 | 0.08 | 0.75 | 0.28 | 0.65 | 1.71 | 0.41 | 0.6 | NA | NA | 0.86 | 0.6 | 0 | 56.23 | 0.39 | NA | 4.92 | 1.82 | 1.89 | 0.73 | 6.35 | NA | 0.23 | 0.26 | 0.81 | 0.44 | 0.67 | 0.21 | 0.7
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 240b (Certified Value 2.99 wt.%) | [32 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.03 | 2.75 | 2.7888 | 3.3147 | 2.9367 | 2.76 | 3.175 | 3.27 | 2.8877 | 2.84 | 3.13 | 3.01 | 2.97 | 3.02 | 2.94 | 2.98 | 3.3759 | 3.375 | 2.89 | 3.24 | 3.4927 | 2.5722 | 2.611 | 2.91 | 2.8 | 3.25 | 2.95 | 2.72 |  |  |  | 1
| 1 | 2 | 3.06 | 2.793 | 2.7169 | 3.2945 | 2.9617 | 2.93 | 3.244 | 3.24 | 2.8898 | 2.84 | 3.19 | 2.95 | 2.9 | 2.97 | 3.01 | 2.88 | 3.2402 | 3.259 | 2.89 | 3.3 | 3.6211 | 2.7916 | 2.545 | 2.92 | 2.85 | 3.37 | 2.9 | 2.74 |  |  |  | #N/A
| 1 | 3 | 3.01 | 2.816 | 2.753 | 3.3045 | 2.8911 | 2.77 | 3.244 | 3.34 | 2.7552 | 2.84 | 3.08 | 3.01 | 2.91 | 3 | 3.01 | 3.08 | 3.319 | 3.247 | 2.79 | 3.19 | 3.9286 | 2.7559 | 2.639 | 2.93 | 2.82 | 3.27 | 2.99 | 2.83 |  |  |  | 16
| 1 | 4 | 3.02 | 2.824 | 2.7301 | 3.2994 | 2.9736 | 2.67 | 3.228 | 3.39 | 2.8051 | 2.84 | 3.11 | 2.96 | 2.99 | 2.98 | 3.03 | 3.12 | 3.2855 | 3.323 | 2.55 | 3.23 | 2.9533 | 2.5467 | 2.5955 | 2.93 | 2.78 | 3.27 | 2.9 | 2.77 |  |  |  | 2.9899
| 1 | 5 | 2.94 | 2.824 | 2.7745 | 3.2909 | 2.9059 | 2.8 | 3.197 | 3.29 | 2.822 | 2.84 | 3.1 | 3.02 | 2.99 | 2.97 | 3.05 | 3.04 | 3.1742 | 3.361 | 2.98 | 3.25 | 3.0697 | 2.6659 | 2.513 | 2.93 | 2.76 | 3.24 | 2.96 | 2.75 |  |  |  | 76
| 1 | 6 | 2.99 | 2.778 | 2.6879 | 3.3159 | 2.8831 | 2.78 | 3.292 | 3.31 | 2.769 | 2.82 | 3.11 | 2.99 | 3.01 | 3 | 3.02 | 2.98 | 3.109 | 3.216 | 2.84 | 3.25 | 3.3875 | 2.6003 | 2.532 | 2.94 | 2.72 | 3.26 | 2.95 | 2.75
| Mean |  | 3.0083 | 2.7975 | 2.7418 | 3.3033 | 2.9253 | 2.785 | 3.23 | 3.3067 | 2.8215 | 2.8367 | 3.12 | 2.99 | 2.9617 | 2.99 | 3.01 | 3.0133 | 3.2506 | 3.2968 | 2.8233 | 3.2433 | 3.4088 | 2.6554 | 2.5726 | 2.9267 | 2.7883 | 3.2767 | 2.9417 | 2.76
| Median |  | 3.015 | 2.8045 | 2.7416 | 3.302 | 2.9213 | 2.775 | 3.236 | 3.3 | 2.8136 | 2.84 | 3.11 | 3 | 2.98 | 2.99 | 3.015 | 3.01 | 3.2628 | 3.291 | 2.865 | 3.245 | 3.4401 | 2.6331 | 2.5702 | 2.93 | 2.79 | 3.265 | 2.95 | 2.75
| Std Dev. |  | 0.040702 | 0.029677 | 0.037593 | 0.010359 | 0.037706 | 0.084083 | 0.040875 | 0.053166 | 0.057391 | 0.008165 | 0.037947 | 0.028983 | 0.045789 | 0.02 | 0.037417 | 0.085479 | 0.097574 | 0.065362 | 0.14801 | 0.03559 | 0.35923 | 0.10053 | 0.04974 | 0.010328 | 0.045789 | 0.047188 | 0.035449 | 0.037947
| Rel.Std.Dev. |  | 1.353% | 1.061% | 1.371% | 0.3136% | 1.289% | 3.019% | 1.265% | 1.608% | 2.034% | 0.2878% | 1.216% | 0.9693% | 1.546% | 0.6689% | 1.243% | 2.837% | 3.002% | 1.983% | 5.242% | 1.097% | 10.54% | 3.786% | 1.933% | 0.3529% | 1.642% | 1.44% | 1.205% | 1.375%
| PDM3 |  | 0.616% | -6.435% | -8.297% | 10.48% | -2.16% | -6.854% | 8.03% | 10.59% | -5.634% | -5.125% | 4.351% | 0.002867% | -0.9448% | 0.002867% | 0.6718% | 0.7833% | 8.72% | 10.27% | -5.571% | 8.476% | 14.01% | -11.19% | -13.96% | -2.115% | -6.742% | 9.591% | -1.614% | -7.69%
| Z-Score (Absolute) |  | 0.12 | 0.66 | 0.86 | 1.21 | 0.19 | 0.7 | 0.94 | 1.22 | 0.57 | 0.51 | 0.53 | 0.05 | 0.05 | 0.05 | 0.13 | 0.14 | 1.01 | 1.18 | 0.56 | 0.99 | 1.6 | 1.18 | 1.49 | 0.18 | 0.69 | 1.11 | 0.13 | 0.8
| Analytical results for As in OREAS 240b (Certified Value 96 ppm) | [32 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 1
| 1 | 1 | 102 | 96.75 | 88.63 | 89.2624 | 96.7964 | 94.5 | 101 | 99 | 94 | 88 | 93.7 | 100 | 97.5 | 98.7 | < 100 | 94.6 | 90 | 86.9521 | 98 | 95.5 | 95 | 68.12 | 92.6716 | 90 | 103.504 | 96.9 | 91.136 | 99.9 | 96 | 96 | 94.7 | 1
| 1 | 2 | 105 | 95.46 | 84.27 | 88.219 | 95.3932 | 97.2 | 103 | 96 | 97 | 89 | 94.8 | 99.7 | 99.9 | 98 | < 100 | 98.4 | 87 | 81.8286 | 93 | 97.2 | 97 | 67.6 | 94.568 | 90 | 104.2893 | 97.5 | 92.782 | 100 | 100 | 96.4 | 93.9 | 19
| 1 | 3 | 105 | 95.28 | 86.63 | 90.9074 | 95.3267 | 94.4 | 99.9 | 99 | 93 | 89 | 93.7 | 99.5 | 97.1 | 97.3 | < 100 | 98 | 90 | 87.6678 | 95 | 92.8 | 94 | 68.54 | 94.6784 | 90 | 101.9299 | 97.9 | 93.336 | 102 | 101 | 98.5 | 96.2 | 16
| 1 | 4 | 103 | 96.63 | 82.88 | 89.6478 | 94.5108 | 91.2 | 103 | 99 | 96 | 89 | 93.9 | 101 | 98.4 | 95 | < 100 | 100.5 | 90 | 85.8841 | 95 | 84.5 | 94 | 67.73 | 95.358 | 95 | 96.6529 | 98.6 | 87.361 | 97.8 | 98 | 95.7 | 97.2 | 95.5824
| 1 | 5 | 101 | 95.87 | 86.59 | 88.8206 | 94.1018 | 94 | 97.2 | 96 | 93 | 90 | 99.2 | 99.6 | 103.5 | 97.7 | < 100 | 100 | 90 | 82.0242 | 98 | 99.6 | 95 | 68.61 | 98.4584 | 90 | 103.0296 | 98.4 | 86.281 | 98.7 | 97 | 97.2 | 94.5 | 77
| 1 | 6 | 103 | 95.07 | 87.24 | 89.488 | 94.8461 | 97.3 | 99.9 | 95 | 93 | 88 | 94.3 | 101 | 100 | 97.9 | < 100 | 98.9 | 91 | 80.3782 | 94 | 94 | 95 | 69.69 | 96.14 | 90 | 101.7887 | 98.7 | 92.121 | 99.6 | 97 | 95.3 | 97.2
| Mean |  | 103.1667 | 95.8433 | 86.04 | 89.3909 | 95.1625 | 94.7667 | 100.6667 | 97.3333 | 94.3333 | 88.8333 | 94.9333 | 100.1333 | 99.4 | 97.4333 |  | 98.4 | 89.6667 | 84.1225 | 95.5 | 93.9333 | 95 | 68.3817 | 95.3124 | 90.8333 | 101.8657 | 98 | 90.5028 | 99.6667 | 98.1667 | 96.5167 | 95.6167
| Median |  | 103 | 95.665 | 86.61 | 89.3752 | 95.0864 | 94.45 | 100.45 | 97.5 | 93.5 | 89 | 94.1 | 99.85 | 99.15 | 97.8 |  | 98.65 | 90 | 83.9541 | 95 | 94.75 | 95 | 68.33 | 95.0182 | 90 | 102.4797 | 98.15 | 91.6285 | 99.75 | 97.5 | 96.2 | 95.45
| Std Dev. |  | 1.6021 | 0.70766 | 2.0937 | 0.90447 | 0.93805 | 2.2739 | 2.1997 | 1.8619 | 1.7512 | 0.75277 | 2.1323 | 0.69186 | 2.3375 | 1.277 |  | 2.089 | 1.3663 | 3.0778 | 2.0736 | 5.206 | 1.0954 | 0.76067 | 1.9243 | 2.0412 | 2.7238 | 0.70427 | 2.9643 | 1.4166 | 1.9408 | 1.1686 | 1.4414
| Rel.Std.Dev. |  | 1.553% | 0.7384% | 2.433% | 1.012% | 0.9857% | 2.399% | 2.185% | 1.913% | 1.856% | 0.8474% | 2.246% | 0.6909% | 2.352% | 1.311% |  | 2.123% | 1.524% | 3.659% | 2.171% | 5.542% | 1.153% | 1.112% | 2.019% | 2.247% | 2.674% | 0.7186% | 3.275% | 1.421% | 1.977% | 1.211% | 1.507%
| PDM3 |  | 7.935% | 0.273% | -9.983% | -6.478% | -0.4393% | -0.8534% | 5.319% | 1.832% | -1.307% | -7.061% | -0.6791% | 4.761% | 3.994% | 1.936% |  | 2.948% | -6.189% | -11.99% | -0.0862% | -1.725% | -0.6093% | -28.46% | -0.2825% | -4.969% | 6.574% | 2.529% | -5.314% | 4.273% | 2.704% | 0.9775% | 0.03586%
| Z-Score (Absolute) |  | 1.72 | 0.12 | 2.03 | 1.29 | 0.03 | 0.12 | 1.17 | 0.44 | 0.21 | 1.41 | 0.08 | 1.05 | 0.89 | 0.46 | 9.9 | 0.67 | 1.23 | 2.44 | 0.04 | 0.3 | 0.07 | 5.88 | 0 | 0.98 | 1.43 | 0.59 | 1.05 | 0.95 | 0.62 | 0.26 | 0.07
| Analytical results for B in OREAS 240b (Certified Value 22.6 ppm) | [32 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.11 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 19 | 20 | 34.67 | 20.5946 | 32 | 30 | 20 | 10 | 20 | 20 | 22 | 17 | 53.26 | 10 | 10 | 11 | < 10 | 26 | 15.4 | [12 empty] | 1
| 1 | 2 | 20 | 19 | 34.03 | 21.2778 | 35 | 28 | 20 | 10 | 20 | 20 | 21 | 17 | 55.39 | 10 | 10 | 12 | < 10 | 27 | 15.6 | [12 empty] | 3
| 1 | 3 | 18 | 19 | 33.42 | 21.7524 | 32 | 28 | 20 | 10 | 20 | 20 | 22 | 15 | 76.69 | 10 | 10 | 12 | < 10 | 28 | 14.5 | [12 empty] | 16
| 1 | 4 | 19 | 19 | 40.03 | 21.4601 | 33 | 30 | 20 | 10 | 20 | 20 | 22 | 15 | 56.14 | 10 | 10 | 12 | < 10 | 27 | 15.4 | [12 empty] | 22.5532
| 1 | 5 | 18 | 19 | 35.91 | 21.168 | 34 | 31 | 20 | 10 | 20 | 20 | 23 | 17 | 55.7 | 10 | 10 | 10 | < 10 | 28 | 14.3 | [12 empty] | 78
| 1 | 6 | 18 | 19 | 33.29 | 21.3714 | 35 | 29 | 20 | 10 | 20 | 20 | 23 | 17 | 64.78 | 10 | 10 | 11 | < 10 | 26 | 15.1
| Mean |  | 18.6667 | 19.1667 | 35.225 | 21.2707 | 33.5 | 29.3333 | 20 | 10 | 20 | 20 | 22.1667 | 16.3333 | 60.3267 | 10 | 10 | 11.3333 |  | 27 | 15.05
| Median |  | 18.5 | 19 | 34.35 | 21.3246 | 33.5 | 29.5 | 20 | 10 | 20 | 20 | 22 | 17 | 55.92 | 10 | 10 | 11.5 |  | 27 | 15.25
| Std Dev. |  | 0.8165 | 0.40825 | 2.5415 | 0.38623 | 1.3784 | 1.2111 | 0 | 0 | 0 | 0 | 0.75277 | 1.0328 | 8.9536 | 0 | 0 | 0.8165 |  | 0.89443 | 0.53198
| Rel.Std.Dev. |  | 4.374% | 2.13% | 7.215% | 1.816% | 4.115% | 4.129% | 0% | 0% | 0% | 0% | 3.396% | 6.323% | 14.84% | 0% | 0% | 7.204% |  | 3.313% | 3.535%
| PDM3 |  | -17.23% | -15.02% | 56.19% | -5.686% | 48.54% | 30.06% | -11.32% | -55.66% | -11.32% | -11.32% | -1.714% | -27.58% | 167.5% | -55.66% | -55.66% | -49.75% |  | 19.72% | -33.27%
| Z-Score (Absolute) |  | 0.28 | 0.23 | 1.51 | 0 | 1.33 | 0.87 | NA | NA | NA | NA | 0.1 | 0.54 | 4.23 | NA | NA | 1.08 | 1.76 | 0.62 | 0.67
Results from laboratories 1.14, 1.15, 1.16, 1.18, 1.27 and 1.29 were removed due to their 10 ppm reading resolution.
| Analytical results for Ba in OREAS 240b (Certified Value 51 ppm) | [32 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 50 | 52 | 52.99 | 57.01 | 52.1935 | 48.3 | 58 | 56 | 50 | 49 | 53 | 50 | 50 | 50 | 68.34 | 50 | 35 | 56.451 | 54 | 43.9 | 50 | 45.46 | 47.088 | 42 | 50 | 49 | 54 | 48 | 46 |  |  | 1
| 1 | 2 | 51 | 52.02 | 52.3 | 56.92 | 51.9114 | 50.1 | 60 | 58 | 52 | 49 | 53 | 50 | 50 | 50 | 66.709 | 50 | 34 | 53.955 | 53 | 42.3 | 51 | 45.73 | 51.156 | 42 | 50 | 50 | 53 | 47 | 46 |  |  | 20
| 1 | 3 | 50 | 52.04 | 51.03 | 56.14 | 52.0289 | 49.6 | 62 | 57 | 49 | 49 | 52 | 50 | 50 | 50 | 64.368 | 50 | 36 | 55.3965 | 53 | 41.7 | 49 | 46.7 | 50.202 | 43 | 50 | 49 | 57 | 48 | 47 |  |  | 16
| 1 | 4 | 50 | 52.13 | 52.33 | 55.77 | 51.6715 | 46.2 | 60 | 58 | 50 | 49 | 52 | 50 | 50 | 50 | 65.664 | 50 | 36 | 55.0498 | 53 | 39.5 | 50 | 46.12 | 46.611 | 42 | 50 | 49 | 56 | 47 | 47 |  |  | 50.6824
| 1 | 5 | 49 | 52 | 53.25 | 56.84 | 50.5066 | 50.4 | 58 | 56 | 51 | 49 | 55 | 50 | 50 | 50 | 64.335 | 50 | 36 | 52.4449 | 54 | 45.1 | 50 | 45.94 | 48.816 | 40 | 50 | 49 | 56 | 48 | 46 |  |  | 79
| 1 | 6 | 50 | 52.02 | 52.01 | 56.01 | 51.7881 | 51.2 | 60 | 58 | 50 | 48 | 53 | 50 | 50 | 50 | 65.601 | 50 | 36 | 51.6335 | 52 | 42.5 | 50 | 46.79 | 47.826 | 40 | 50 | 49 | 55 | 47 | 46
| Mean |  | 50 | 52.035 | 52.3183 | 56.4483 | 51.6833 | 49.3 | 59.6667 | 57.1667 | 50.3333 | 48.8333 | 53 | 50 | 50 | 50 | 65.8362 | 50 | 35.5 | 54.1551 | 53.1667 | 42.5 | 50 | 46.1233 | 48.6165 | 41.5 | 50 | 49.1667 | 55.1667 | 47.5 | 46.3333
| Median |  | 50 | 52.02 | 52.315 | 56.49 | 51.8498 | 49.85 | 60 | 57.5 | 50 | 49 | 53 | 50 | 50 | 50 | 65.6325 | 50 | 36 | 54.5024 | 53 | 42.4 | 50 | 46.03 | 48.321 | 42 | 50 | 49 | 55.5 | 47.5 | 46
| Std Dev. |  | 0.63246 | 0.048888 | 0.78415 | 0.53641 | 0.60458 | 1.7978 | 1.5055 | 0.98319 | 1.0328 | 0.40825 | 1.0954 | 0 | 0 | 0 | 1.5182 | 0 | 0.83666 | 1.8404 | 0.75277 | 1.9183 | 0.63246 | 0.53016 | 1.7881 | 1.2247 | 0 | 0.40825 | 1.472 | 0.54772 | 0.5164
| Rel.Std.Dev. |  | 1.265% | 0.09395% | 1.499% | 0.9503% | 1.17% | 3.647% | 2.523% | 1.72% | 2.052% | 0.836% | 2.067% | 0% | 0% | 0% | 2.306% | 0% | 2.357% | 3.398% | 1.416% | 4.514% | 1.265% | 1.149% | 3.678% | 2.951% | 0% | 0.8303% | 2.668% | 1.153% | 1.115%
| PDM3 |  | -1.346% | 2.669% | 3.228% | 11.38% | 1.975% | -2.728% | 17.73% | 12.79% | -0.6888% | -3.648% | 4.573% | -1.346% | -1.346% | -1.346% | 29.9% | -1.346% | -29.96% | 6.852% | 4.902% | -16.14% | -1.346% | -8.995% | -4.076% | -18.12% | -1.346% | -2.991% | 8.848% | -6.279% | -8.581%
| Z-Score (Absolute) |  | 0.04 | 0.43 | 0.5 | 1.45 | 0.35 | 0.2 | 2.2 | 1.62 | 0.04 | 0.31 | 0.66 | NA | NA | NA | 3.62 | NA | 3.39 | 0.92 | 0.69 | 1.77 | 0.04 | 0.93 | 0.36 | 2 | NA | 0.23 | 1.16 | 0.62 | 0.89
| Analytical results for Be in OREAS 240b (Certified Value 0.25 ppm) | [32 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.3 | 0.29 | < 1 | 0.17025 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.27 | < 5 | < 1 | 0.25 | 0.24 | 0.25 | < 10 | 0.24 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.12 | 0.3 | 0.3 | 0.27 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.24 | < 1 | 0.18068 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.31 | < 5 | < 1 | 0.25 | 0.25 | 0.25 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.09 | 0.3 | 0.29 | 0.18 | [5 empty] | 21
| 1 | 3 | 0.2 | 0.28 | < 1 | 0.17905 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.29 | < 5 | < 1 | 0.25 | 0.25 | 0.24 | < 10 | 0.26 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.14 | 0.3 | 0.32 | 0.2 | [5 empty] | 16
| 1 | 4 | 0.3 | 0.3 | < 1 | 0.16589 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.32 | < 5 | < 1 | 0.25 | 0.26 | 0.24 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.11 | 0.3 | 0.28 | 0.14 | [5 empty] | 0.24784
| 1 | 5 | 0.2 | 0.28 | < 1 | 0.16688 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.27 | < 5 | < 1 | 0.25 | 0.25 | 0.24 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.1 | 0.3 | 0.3 | 0.15 | [5 empty] | 80
| 1 | 6 | 0.2 | 0.28 | < 1 | 0.17139 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.28 | < 5 | < 1 | 0.25 | 0.24 | 0.25 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.07 | 0.3 | 0.3 | 0.25
| Mean |  | 0.25 | 0.27833 |  | 0.17236 |  | 0.3 |  |  | 0.29 |  |  | 0.25 | 0.24833 | 0.245 |  | 0.25 |  |  |  |  | 0.2 | 0.24 | 0.105 | 0.3 | 0.29833 | 0.19833
| Median |  | 0.25 | 0.28 |  | 0.17082 |  | 0.3 |  |  | 0.285 |  |  | 0.25 | 0.25 | 0.245 |  | 0.25 |  |  |  |  | 0.2 | 0.24 | 0.105 | 0.3 | 0.3 | 0.19
| Std Dev. |  | 0.054772 | 0.020412 |  | 0.0061874 |  | 0 |  |  | 0.020976 |  |  | 0 | 0.0075277 | 0.0054772 |  | 0.0063246 |  |  |  |  | 3.0405e-17 | 0 | 0.02429 | 0 | 0.013292 | 0.052694
| Rel.Std.Dev. |  | 21.91% | 7.334% |  | 3.59% |  | 0% |  |  | 7.233% |  |  | 0% | 3.031% | 2.236% |  | 2.53% |  |  |  |  | 1.52e-14% | 0% | 23.13% | 0% | 4.455% | 26.57%
| PDM3 |  | 0.8733% | 12.31% |  | -30.46% |  | 21.05% |  |  | 17.01% |  |  | 0.8733% | 0.2008% | -1.144% |  | 0.8733% |  |  |  |  | -19.3% | -3.162% | -57.63% | 21.05% | 20.38% | -19.97%
| Z-Score (Absolute) |  | NA | 1 | 8.8 | 2.73 | 0 | NA | 0 | 0 | 1.41 | 79.16 | 8.8 | 0 | 0.06 | 0.18 | 167.11 | 0 | 0 | 8.8 | 0 | 0 | NA | 0.35 | 5.1 | NA | 1.7 | 1.82
Results from laboratories 1.01, 1.07, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 240b (Certified Value 0.067 ppm) | [32 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.05 | 0.07 | < 1 | 0.052628 | 0.07 | 0.08 | 0.05 | 0.09 | < 2 | < 0.1 | 0.07 | 0.07 | 0.07 | < 10 | 0.06 | < 5 | 0.06 | 0.06 | 0.07 | < 0.1 | 0.07517 | 0.06 | 0.057 | 0.1 | 0.07 | 0.08 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.05 | 0.07 | < 1 | 0.048489 | 0.07 | 0.1 | 0.05 | 0.07 | < 2 | < 0.1 | 0.06 | 0.06 | 0.08 | < 10 | 0.07 | < 5 | 0.06 | 0.05 | 0.07 | < 0.1 | 0.07575 | 0.07 | 0.051 | 0.09 | 0.07 | 0.07 | 0.06 |  |  |  |  | 22
| 1 | 3 | 0.05 | 0.07 | < 1 | 0.050028 | 0.09 | 0.09 | < 0.05 | 0.07 | < 2 | < 0.1 | 0.07 | 0.07 | 0.07 | < 10 | 0.07 | < 5 | 0.07 | 0.05 | 0.07 | < 0.1 | 0.07735 | 0.06 | 0.054 | 0.08 | 0.07 | 0.08 | 0.06 |  |  |  |  | 16
| 1 | 4 | 0.06 | 0.07 | < 1 | 0.054331 | 0.07 | 0.08 | 0.05 | 0.07 | < 2 | < 0.1 | 0.07 | 0.07 | 0.07 | < 10 | 0.07 | < 5 | 0.07 | 0.05 | 0.07 | < 0.1 | 0.08203 | 0.06 | 0.052 | 0.08 | 0.07 | 0.07 | 0.06 |  |  |  |  | 0.067386
| 1 | 5 | 0.05 | 0.07 | < 1 | 0.043873 | 0.07 | 0.08 | 0.05 | 0.07 | < 2 | < 0.1 | 0.06 | 0.07 | 0.07 | < 10 | 0.07 | < 5 | 0.07 | 0.06 | 0.07 | < 0.1 | 0.07216 | 0.06 | 0.051 | 0.08 | 0.08 | 0.07 | 0.05 |  |  |  |  | 81
| 1 | 6 | 0.05 | 0.07 | < 1 | 0.048437 | 0.07 | 0.08 | 0.07 | 0.06 | < 2 | < 0.1 | 0.06 | 0.07 | 0.07 | < 10 | 0.07 | < 5 | 0.07 | 0.06 | 0.07 | < 0.1 | 0.0808 | 0.07 | 0.06 | 0.08 | 0.07 | 0.08 | 0.06
| Mean |  | 0.051667 | 0.07 |  | 0.049631 | 0.073333 | 0.085 | 0.054 | 0.071667 |  |  | 0.065 | 0.068333 | 0.071667 |  | 0.068333 |  | 0.066667 | 0.055 | 0.07 |  | 0.07721 | 0.063333 | 0.054167 | 0.085 | 0.071667 | 0.075 | 0.056667
| Median |  | 0.05 | 0.07 |  | 0.049258 | 0.07 | 0.08 | 0.05 | 0.07 |  |  | 0.065 | 0.07 | 0.07 |  | 0.07 |  | 0.07 | 0.055 | 0.07 |  | 0.07655 | 0.06 | 0.053 | 0.08 | 0.07 | 0.075 | 0.06
| Std Dev. |  | 0.0040825 | 0 |  | 0.0036636 | 0.008165 | 0.0083666 | 0.0089443 | 0.0098319 |  |  | 0.0054772 | 0.0040825 | 0.0040825 |  | 0.0040825 |  | 0.005164 | 0.0054772 | 0 |  | 0.0036861 | 0.005164 | 0.003656 | 0.0083666 | 0.0040825 | 0.0054772 | 0.005164
| Rel.Std.Dev. |  | 7.902% | 0% |  | 7.382% | 11.13% | 9.843% | 16.56% | 13.72% |  |  | 8.427% | 5.974% | 5.696% |  | 5.974% |  | 7.746% | 9.959% | 0% |  | 4.774% | 8.154% | 6.75% | 9.843% | 5.696% | 7.303% | 9.113%
| PDM3 |  | -23.33% | 3.88% |  | -26.35% | 8.827% | 26.14% | -19.86% | 6.353% |  |  | -3.54% | 1.407% | 6.353% |  | 1.407% |  | -1.067% | -18.38% | 3.88% |  | 14.58% | -6.013% | -19.62% | 26.14% | 6.353% | 11.3% | -15.91%
| Z-Score (Absolute) |  | 0.93 | 0 | 21.75 | 1.03 | 0.17 | 0.76 | 1.05 | 0.08 | 47.03 | 1.01 | 0.25 | 0.08 | 0.08 | 249.33 | 0.08 | 122.89 | 0.17 | 0.76 | 0 | 1.01 | 0.36 | 0.34 | 0.8 | 0.76 | 0.08 | 0.25 | 0.67
| Analytical results for Ca in OREAS 240b (Certified Value 2.24 wt.%) | [32 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 2.23 | 2.208 | 2.4937 | 2.5988 | 2.3258 | 2.73 | 2.423 | 2.75 | 2.41 | 2.34 | 2.02 | 2.17 | 2.13 | 2.13 | 1.8291 | 2.09 | 2.06 | 2.6935 | 2.42 | 2.1 | 2.19 | 2.1287 | 2.3485 | 1.5805 | 2.16 | 1.93 | 2.08 | 2.18 | 1.82 |  |  | 1
| 1 | 2 | 2.24 | 2.177 | 2.4243 | 2.5554 | 2.2766 | 2.82 | 2.594 | 2.7 | 2.49 | 2.38 | 2.05 | 2.13 | 2.06 | 2.11 | 1.7675 | 2.18 | 1.97 | 2.5895 | 2.6 | 2.1 | 2.24 | 2.136 | 2.4235 | 1.534 | 2.16 | 1.98 | 2.16 | 2.15 | 1.83 |  |  | #N/A
| 1 | 3 | 2.16 | 2.187 | 2.4267 | 2.5912 | 2.2796 | 2.71 | 2.573 | 2.79 | 2.33 | 2.35 | 1.93 | 2.17 | 2.07 | 2.12 | 1.7156 | 2.2 | 2.13 | 2.7049 | 2.6 | 1.98 | 2.17 | 2.31 | 2.3871 | 1.575 | 2.15 | 1.96 | 2.09 | 2.21 | 1.86 |  |  | 16
| 1 | 4 | 2.18 | 2.191 | 2.4402 | 2.5786 | 2.2774 | 2.68 | 2.537 | 2.86 | 2.45 | 2.34 | 1.95 | 2.14 | 2.19 | 2.11 | 1.7386 | 2.2 | 2.18 | 2.5346 | 2.41 | 1.84 | 2.19 | 1.7926 | 2.2859 | 1.5635 | 2.16 | 1.91 | 2.11 | 2.15 | 1.83 |  |  | 2.2395
| 1 | 5 | 2.11 | 2.198 | 2.5445 | 2.5591 | 2.2293 | 2.74 | 2.473 | 2.77 | 2.37 | 2.34 | 1.95 | 2.16 | 2.19 | 2.09 | 1.729 | 2.24 | 2.15 | 2.5415 | 2.37 | 2.12 | 2.19 | 1.8695 | 2.2726 | 1.4955 | 2.17 | 1.83 | 2.08 | 2.21 | 1.81 |  |  | 82
| 1 | 6 | 2.16 | 2.178 | 2.3727 | 2.5816 | 2.3055 | 2.75 | 2.616 | 2.79 | 2.37 | 2.36 | 1.97 | 2.14 | 2.2 | 2.12 | 1.7312 | 2.18 | 2.07 | 2.4736 | 2.36 | 2.07 | 2.19 | 1.7926 | 2.2518 | 1.502 | 2.19 | 1.8 | 2.09 | 2.17 | 1.77
| Mean |  | 2.18 | 2.1898 | 2.4503 | 2.5774 | 2.2824 | 2.7383 | 2.536 | 2.7767 | 2.4033 | 2.3517 | 1.9783 | 2.1517 | 2.14 | 2.1133 | 1.7518 | 2.1817 | 2.0933 | 2.5896 | 2.46 | 2.035 | 2.195 | 2.0049 | 2.3282 | 1.5418 | 2.165 | 1.9017 | 2.1017 | 2.1783 | 1.82
| Median |  | 2.17 | 2.189 | 2.4334 | 2.5801 | 2.2785 | 2.735 | 2.555 | 2.78 | 2.39 | 2.345 | 1.96 | 2.15 | 2.16 | 2.115 | 1.7349 | 2.19 | 2.1 | 2.5655 | 2.415 | 2.085 | 2.19 | 1.9991 | 2.3172 | 1.5488 | 2.16 | 1.92 | 2.09 | 2.175 | 1.825
| Std Dev. |  | 0.04858 | 0.011923 | 0.060204 | 0.017242 | 0.032576 | 0.047081 | 0.074662 | 0.052789 | 0.058878 | 0.016021 | 0.046655 | 0.017224 | 0.063246 | 0.013663 | 0.041589 | 0.049967 | 0.07607 | 0.092621 | 0.11082 | 0.10766 | 0.023452 | 0.21638 | 0.06882 | 0.037039 | 0.013784 | 0.071949 | 0.030605 | 0.027142 | 0.029665
| Rel.Std.Dev. |  | 2.228% | 0.5445% | 2.457% | 0.6689% | 1.427% | 1.719% | 2.944% | 1.901% | 2.45% | 0.6813% | 2.358% | 0.8005% | 2.955% | 0.6465% | 2.374% | 2.29% | 3.634% | 3.577% | 4.505% | 5.29% | 1.068% | 10.79% | 2.956% | 2.402% | 0.6367% | 3.783% | 1.456% | 1.246% | 1.63%
| PDM3 |  | -2.656% | -2.217% | 9.415% | 15.09% | 1.915% | 22.27% | 13.24% | 23.99% | 7.316% | 5.009% | -11.66% | -3.922% | -4.443% | -5.633% | -21.77% | -2.582% | -6.526% | 15.63% | 9.846% | -9.131% | -1.987% | -10.48% | 3.963% | -31.16% | -3.326% | -15.08% | -6.154% | -2.731% | -18.73%
| Z-Score (Absolute) |  | 0 | 0.04 | 1.06 | 1.56 | 0.4 | 2.19 | 1.4 | 2.34 | 0.88 | 0.67 | 0.79 | 0.11 | 0.16 | 0.26 | 1.68 | 0.01 | 0.34 | 1.61 | 1.1 | 0.57 | 0.06 | 0.69 | 0.58 | 2.51 | 0.06 | 1.09 | 0.31 | 0.01 | 1.41
| Analytical results for Cd in OREAS 240b (Certified Value 0.86 ppm) | [32 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.84 | 0.862 | 1.13 | 0.64067 | 0.8801 | 0.84 | 0.95 | 0.92 | 0.83 | 1 | 0.9 | 0.86 | 0.88 | 0.85 | < 10 | 0.86 | < 1 | < 1 | 0.78 | 0.74 | 0.82 | 1.78 | 1 | 0.88066 | 0.87 | 0.84 | 0.86 | 0.8 | 0.86 |  |  | 1
| 1 | 2 | 0.89 | 0.866 | 1.11 | 0.67516 | 0.86631 | 0.94 | 1.05 | 0.91 | 0.89 | 1 | 0.9 | 0.85 | 0.89 | 0.82 | < 10 | 0.92 | < 1 | < 1 | 0.74 | 0.77 | 0.86 | 1.75 | 1 | 0.82556 | 0.87 | 0.8 | 0.85 | 0.79 | 0.84 |  |  | 23
| 1 | 3 | 0.98 | 0.859 | 1.13 | 0.65873 | 0.91334 | 0.87 | 0.94 | 0.9 | 0.85 | 1 | 0.9 | 0.82 | 0.88 | 0.85 | < 10 | 0.89 | < 1 | < 1 | 0.84 | 0.74 | 0.87 | 1.77 | 1 | 0.85863 | 0.88 | 0.78 | 0.88 | 0.8 | 0.86 |  |  | 16
| 1 | 4 | 0.97 | 0.862 | 1.09 | 0.63861 | 0.88052 | 0.81 | 0.98 | 0.92 | 0.9 | 1 | 0.9 | 0.87 | 0.89 | 0.81 | < 10 | 0.91 | < 1 | < 1 | 0.84 | 0.65 | 0.9 | 1.76 | 0.5 | 0.84188 | 0.86 | 0.81 | 0.87 | 0.79 | 0.91 |  |  | 0.85933
| 1 | 5 | 0.9 | 0.839 | 1.08 | 0.63928 | 0.8676 | 0.82 | 0.94 | 0.97 | 0.88 | 1 | 0.9 | 0.84 | 0.89 | 0.83 | < 10 | 0.93 | < 1 | < 1 | 0.72 | 0.76 | 0.89 | 1.71 | 0.5 | 0.94378 | 0.89 | 0.81 | 0.88 | 0.82 | 0.81 |  |  | 83
| 1 | 6 | 0.86 | 0.843 | 1.11 | 0.64518 | 0.8521 | 0.95 | 0.97 | 0.96 | 0.82 | 1 | 0.9 | 0.85 | 0.89 | 0.8 | < 10 | 0.95 | < 1 | < 1 | 0.77 | 0.73 | 0.87 | 1.74 | 0.5 | 0.73411 | 0.88 | 0.78 | 0.89 | 0.8 | 0.89
| Mean |  | 0.90667 | 0.85517 | 1.1083 | 0.64961 | 0.87666 | 0.87167 | 0.97167 | 0.93 | 0.86167 | 1 | 0.9 | 0.84833 | 0.88667 | 0.82667 |  | 0.91 |  |  | 0.78167 | 0.73167 | 0.86833 | 1.7517 | 0.75 | 0.84744 | 0.875 | 0.80333 | 0.87167 | 0.8 | 0.86167
| Median |  | 0.895 | 0.8605 | 1.11 | 0.64293 | 0.87385 | 0.855 | 0.96 | 0.92 | 0.865 | 1 | 0.9 | 0.85 | 0.89 | 0.825 |  | 0.915 |  |  | 0.775 | 0.74 | 0.87 | 1.755 | 0.75 | 0.85025 | 0.875 | 0.805 | 0.875 | 0.8 | 0.86
| Std Dev. |  | 0.057155 | 0.011268 | 0.020412 | 0.014583 | 0.020808 | 0.06047 | 0.041673 | 0.028284 | 0.033116 | 0 | 0 | 0.017224 | 0.005164 | 0.020656 |  | 0.031623 |  |  | 0.049967 | 0.042622 | 0.027869 | 0.024833 | 0.27386 | 0.069085 | 0.010488 | 0.022509 | 0.01472 | 0.010954 | 0.035449
| Rel.Std.Dev. |  | 6.304% | 1.318% | 1.842% | 2.245% | 2.373% | 6.937% | 4.289% | 3.041% | 3.843% | 0% | 0% | 2.03% | 0.5824% | 2.499% |  | 3.475% |  |  | 6.392% | 5.825% | 3.209% | 1.418% | 36.51% | 8.152% | 1.199% | 2.802% | 1.689% | 1.369% | 4.114%
| PDM3 |  | 5.509% | -0.4845% | 28.98% | -24.41% | 2.017% | 1.436% | 13.07% | 8.224% | 0.2719% | 16.37% | 4.733% | -1.28% | 3.181% | -3.801% |  | 5.896% |  |  | -9.038% | -14.86% | 1.048% | 103.8% | -12.72% | -1.384% | 1.824% | -6.516% | 1.436% | -6.904% | 0.2719%
| Z-Score (Absolute) |  | 0.65 | 0.15 | 3.79 | 3.35 | 0.18 | 0.1 | 1.66 | 1.01 | 0.05 | NA | NA | 0.26 | 0.34 | 0.6 | 64.34 | 0.7 | 5.68 | 5.68 | 1.3 | 2.07 | 0.05 | 13.8 | NA | 0.27 | 0.16 | 0.96 | 0.1 | 1.01 | 0.05
Results from laboratories 1.12, 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 240b (Certified Value 11.7 ppm) | [32 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [8 empty] | 2
| 1 | 1 | 11.02 | 11.968 | 10.479 | 9.3702 | 12.0315 | 11.6 | 4.16 | 13.7 | 11.423 | 13 | 11.1 | 11.9 | 11.65 | 10.85 | < 10 | 12.05 | 11.96 | 11.5 | 12.4 | 12.7 | 11.34 | 12.06 | 11.8 | [8 empty] | 1
| 1 | 2 | 11.26 | 11.876 | 10.504 | 9.3306 | 11.8658 | 12.1 | 3.78 | 13.3 | 11.59 | 13 | 11.55 | 12 | 11.6 | 11.45 | < 10 | 11.15 | 12.4 | 11 | 12.3 | 12.8 | 11.5 | 12.06 | 11.5 | [8 empty] | 24
| 1 | 3 | 11.13 | 11.892 | 10.672 | 9.3982 | 11.9628 | 11.6 | 3.77 | 13.3 | 11.3 | 13 | 10.85 | 12 | 12 | 11.2 | < 10 | 11.43 | 12 | 11 | 12.4 | 12.9 | 11.85 | 12.3 | 11.7 | [8 empty] | 16
| 1 | 4 | 11.08 | 12.106 | 10.521 | 9.48 | 11.7474 | 11.5 | 3.43 | 13.7 | 11.282 | 13 | 11.3 | 12.15 | 11.05 | 11.7 | < 10 | 12.09 | 12.15 | 11.5 | 12.4 | 12.7 | 11.72 | 11.97 | 11.9 | [8 empty] | 11.6936
| 1 | 5 | 10.94 | 11.981 | 10.935 | 9.3858 | 11.7147 | 11.9 | 4.55 | 13.3 | 11.424 | 13 | 11.95 | 12 | 11.4 | 11.65 | < 10 | 11.52 | 12.16 | 10.5 | 12.55 | 12.4 | 11.86 | 12.32 | 11.5 | [8 empty] | 84
| 1 | 6 | 11.11 | 11.756 | 11.522 | 9.4041 | 12.0257 | 12.2 | 3.73 | 13.4 | 11.064 | 13 | 12.2 | 12 | 11.35 | 11.65 | < 10 | 11.58 | 12.38 | 10.5 | 12.45 | 12.4 | 11.73 | 12 | 11.8
| Mean |  | 11.09 | 11.9298 | 10.7722 | 9.3948 | 11.8913 | 11.8167 | 3.9033 | 13.45 | 11.3472 | 13 | 11.4917 | 12.0083 | 11.5083 | 11.4167 |  | 11.6367 | 12.175 | 11 | 12.4167 | 12.65 | 11.6667 | 12.1183 | 11.7
| Median |  | 11.095 | 11.93 | 10.5965 | 9.392 | 11.9143 | 11.75 | 3.775 | 13.35 | 11.3615 | 13 | 11.425 | 12 | 11.5 | 11.55 |  | 11.55 | 12.155 | 11 | 12.4 | 12.7 | 11.725 | 12.06 | 11.75
| Std Dev. |  | 0.10807 | 0.11803 | 0.40491 | 0.049313 | 0.13813 | 0.29269 | 0.39272 | 0.19748 | 0.17736 | 0 | 0.51324 | 0.080104 | 0.32158 | 0.33417 |  | 0.36681 | 0.18458 | 0.44721 | 0.08165 | 0.20736 | 0.20607 | 0.15263 | 0.16733
| Rel.Std.Dev. |  | 0.9745% | 0.9894% | 3.759% | 0.5249% | 1.162% | 2.477% | 10.06% | 1.468% | 1.563% | 0% | 4.466% | 0.6671% | 2.794% | 2.927% |  | 3.152% | 1.516% | 4.066% | 0.6576% | 1.639% | 1.766% | 1.26% | 1.43%
| PDM3 |  | -5.162% | 2.02% | -7.88% | -19.66% | 1.69% | 1.052% | -66.62% | 15.02% | -2.963% | 11.17% | -1.727% | 2.691% | -1.585% | -2.369% |  | -0.4872% | 4.116% | -5.932% | 6.183% | 8.178% | -0.2307% | 3.632% | 0.05439%
| Z-Score (Absolute) |  | 0.92 | 0.46 | 1.44 | 3.7 | 0.39 | 0.27 | 12.69 | 2.95 | 0.5 | NA | 0.26 | 0.58 | 0.23 | 0.38 | 10.9 | 0.02 | 0.86 | 1.07 | 1.25 | 1.64 | 0.02 | 0.76 | 0.08
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 240b (Certified Value 29.8 ppm) | [32 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 28.5 | 29.1 | 27.59 | 22.5907 | 30.9112 | 31 | 30.3 | 32.7 | 31 | 27 | 30.3 | 30.1 | 30.4 | 30.2 | 24.546 | 27.5 | 28 | 33.4286 | 31.8 | 29.3 | 30 | 26.07 | 27 | 32.3217 | 28.8 | 33.8 | 29.7 | 29.5 | 29.2 |  |  | 1
| 1 | 2 | 28.9 | 29.65 | 26.97 | 22.5325 | 30.371 | 31.7 | 31.2 | 31.5 | 31.7 | 27 | 30.4 | 30.5 | 30.1 | 30 | 24.331 | 29 | 27 | 32.495 | 32.2 | 28.8 | 31 | 27.95 | 26 | 31.888 | 28.2 | 33.9 | 30.1 | 29.4 | 29 |  |  | 25
| 1 | 3 | 29.2 | 29.81 | 27.54 | 22.3095 | 30.5705 | 30.4 | 30.2 | 32.3 | 29.8 | 27 | 29.9 | 29 | 30 | 30.3 | 24.919 | 28.8 | 28 | 33.2369 | 31.2 | 28 | 30 | 26.22 | 27 | 31.056 | 28 | 35 | 31.9 | 30.1 | 29.7 |  |  | 16
| 1 | 4 | 29.2 | 29.93 | 26.98 | 22.0212 | 30.1216 | 29.6 | 30.1 | 32.8 | 31.2 | 27 | 30.2 | 30.4 | 30.5 | 29 | 25.358 | 29.9 | 28 | 32.946 | 31.5 | 25.5 | 30 | 29.47 | 27 | 30.2656 | 28.7 | 34 | 30.9 | 29.1 | 29.8 |  |  | 29.8154
| 1 | 5 | 28.4 | 29.6 | 27.42 | 22.0229 | 30.0438 | 31 | 29.8 | 31.8 | 30 | 27 | 32 | 30.1 | 30.4 | 29.7 | 23.828 | 29.4 | 28 | 32.198 | 32.4 | 30.1 | 31 | 27.6 | 26 | 32.1086 | 28.8 | 33.6 | 30.7 | 29.9 | 29.1 |  |  | 85
| 1 | 6 | 28.7 | 29.18 | 26.68 | 22.2579 | 30.6811 | 31.6 | 29.2 | 31.2 | 30 | 27 | 30.3 | 30.8 | 30.2 | 30.4 | 24.92 | 29.8 | 28 | 31.3228 | 31.2 | 29.8 | 30 | 29.09 | 25 | 30.7702 | 28.5 | 34.1 | 31.2 | 29.3 | 29.8
| Mean |  | 28.8167 | 29.545 | 27.1967 | 22.2891 | 30.4499 | 30.8833 | 30.1333 | 32.05 | 30.6167 | 27 | 30.5167 | 30.15 | 30.2667 | 29.9333 | 24.6503 | 29.0667 | 27.8333 | 32.6046 | 31.7167 | 28.5833 | 30.3333 | 27.7333 | 26.3333 | 31.4017 | 28.5 | 34.0667 | 30.75 | 29.55 | 29.4333
| Median |  | 28.8 | 29.625 | 27.2 | 22.2837 | 30.4707 | 31 | 30.15 | 32.05 | 30.5 | 27 | 30.3 | 30.25 | 30.3 | 30.1 | 24.7325 | 29.2 | 28 | 32.7205 | 31.65 | 29.05 | 30 | 27.775 | 26.5 | 31.472 | 28.6 | 33.95 | 30.8 | 29.45 | 29.45
| Std Dev. |  | 0.34303 | 0.33578 | 0.37087 | 0.24257 | 0.33458 | 0.78592 | 0.65625 | 0.65345 | 0.78592 | 0 | 0.74677 | 0.62209 | 0.19664 | 0.52026 | 0.53568 | 0.88015 | 0.40825 | 0.77664 | 0.50761 | 1.6845 | 0.5164 | 1.4127 | 0.8165 | 0.82359 | 0.33466 | 0.48854 | 0.78422 | 0.37815 | 0.37238
| Rel.Std.Dev. |  | 1.19% | 1.137% | 1.364% | 1.088% | 1.099% | 2.545% | 2.178% | 2.039% | 2.567% | 0% | 2.447% | 2.063% | 0.6497% | 1.738% | 2.173% | 3.028% | 1.467% | 2.382% | 1.6% | 5.893% | 1.702% | 5.094% | 3.101% | 2.623% | 1.174% | 1.434% | 2.55% | 1.28% | 1.265%
| PDM3 |  | -3.35% | -0.9068% | -8.783% | -25.24% | 2.128% | 3.582% | 1.066% | 7.495% | 2.688% | -9.443% | 2.352% | 1.122% | 1.514% | 0.3956% | -17.32% | -2.511% | -6.648% | 9.355% | 6.377% | -4.132% | 1.737% | -6.983% | -11.68% | 5.32% | -4.412% | 14.26% | 3.135% | -0.8901% | -1.281%
| Z-Score (Absolute) |  | 0.67 | 0.23 | 1.65 | 4.62 | 0.31 | 0.57 | 0.12 | 1.28 | 0.41 | 1.77 | 0.35 | 0.13 | 0.2 | 0 | 3.19 | 0.52 | 1.27 | 1.61 | 1.08 | 0.82 | 0.24 | 1.33 | 2.17 | 0.89 | 0.87 | 2.5 | 0.49 | 0.23 | 0.3
| Analytical results for Cr in OREAS 240b (Certified Value 34.7 ppm) | [32 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 37 | 34.8 | 36.97 | 35.8059 | 40 | 37 | 42 | 39 | 34 | 38 | 33 | 33 | 32 | 31.913 | 32 | 30 | 37.7425 | 36 | 35.2 | 35 | 32.77 | 33.2847 | 30 | 34 | 33.8 | 35 | 32 | 31.2 |  |  |  | 1
| 1 | 2 | 38 | 35 | 36.38 | 35.1868 | 42 | 38 | 40 | 39 | 34 | 38 | 32 | 32 | 32 | 29.975 | 33 | 29 | 36.5778 | 35 | 34.8 | 35 | 39.08 | 34.6518 | 30 | 33 | 34 | 34 | 32 | 31.2 |  |  |  | 26
| 1 | 3 | 37 | 35.4 | 36.55 | 35.6804 | 40 | 39 | 41 | 39 | 34 | 36 | 33 | 33 | 32 | 29.633 | 33 | 31 | 37.8518 | 35 | 33.9 | 34 | 35.57 | 35.0517 | 30 | 34 | 34.3 | 35 | 33 | 32.5 |  |  |  | 16
| 1 | 4 | 37 | 35.4 | 37.18 | 35.4541 | 39 | 39 | 45 | 40 | 35 | 37 | 33 | 33 | 32 | 29.874 | 33 | 31 | 37.376 | 36 | 30.7 | 34 | 34 | 33.0987 | 30 | 34 | 33.2 | 36 | 32 | 32.3 |  |  |  | 34.6791
| 1 | 5 | 36 | 35.1 | 36.59 | 34.7861 | 41 | 38 | 41 | 37 | 34 | 39 | 33 | 34 | 32 | 29.769 | 33 | 31 | 35.6477 | 34 | 36.2 | 35 | 34.48 | 34.8192 | 30 | 34 | 33 | 35 | 32 | 31.7 |  |  |  | 86
| 1 | 6 | 37 | 34.3 | 36.15 | 35.4663 | 41 | 38 | 41 | 36 | 33 | 37 | 33 | 34 | 32 | 30.056 | 33 | 31 | 35.1796 | 34 | 34 | 35 | 36.43 | 34.2054 | 30 | 34 | 32.4 | 35 | 32 | 31.5
| Mean |  | 37 | 35 | 36.6367 | 35.3966 | 40.5 | 38.1667 | 41.6667 | 38.3333 | 34 | 37.5 | 32.8333 | 33.1667 | 32 | 30.2033 | 32.8333 | 30.5 | 36.7292 | 35 | 34.1333 | 34.6667 | 35.3883 | 34.1853 | 30 | 33.8333 | 33.45 | 35 | 32.1667 | 31.7333
| Median |  | 37 | 35.05 | 36.57 | 35.4602 | 40.5 | 38 | 41 | 39 | 34 | 37.5 | 33 | 33 | 32 | 29.9245 | 33 | 31 | 36.9769 | 35 | 34.4 | 35 | 35.025 | 34.4286 | 30 | 34 | 33.5 | 35 | 32 | 31.6
| Std Dev. |  | 0.63246 | 0.41473 | 0.37914 | 0.3669 | 1.0488 | 0.75277 | 1.7512 | 1.5055 | 0.63246 | 1.0488 | 0.40825 | 0.75277 | 0 | 0.85079 | 0.40825 | 0.83666 | 1.1224 | 0.89443 | 1.8822 | 0.5164 | 2.207 | 0.82003 | 0 | 0.40825 | 0.70922 | 0.63246 | 0.40825 | 0.55377
| Rel.Std.Dev. |  | 1.709% | 1.185% | 1.035% | 1.037% | 2.59% | 1.972% | 4.203% | 3.928% | 1.86% | 2.797% | 1.243% | 2.27% | 0% | 2.817% | 1.243% | 2.743% | 3.056% | 2.556% | 5.514% | 1.49% | 6.237% | 2.399% | 0% | 1.207% | 2.12% | 1.807% | 1.269% | 1.745%
| PDM3 |  | 6.693% | 0.9254% | 5.645% | 2.069% | 16.79% | 10.06% | 20.15% | 10.54% | -1.958% | 8.134% | -5.322% | -4.361% | -7.725% | -12.91% | -5.322% | -12.05% | 5.912% | 0.9254% | -1.574% | -0.03578% | 2.045% | -1.424% | -13.49% | -2.439% | -3.544% | 0.9254% | -7.245% | -8.494%
| Z-Score (Absolute) |  | 0.91 | 0.2 | 0.78 | 0.34 | 2.15 | 1.33 | 2.57 | 1.39 | 0.15 | 1.09 | 0.56 | 0.45 | 0.86 | 1.5 | 0.56 | 1.39 | 0.82 | 0.2 | 0.1 | 0.09 | 0.34 | 0.09 | 1.57 | 0.21 | 0.35 | 0.2 | 0.8 | 0.95
| Analytical results for Cs in OREAS 240b (Certified Value 1 ppm) | [32 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | 1.12 | 0.98 | 0.81005 | 0.84 | 0.94 | 1.02 | 1 | 1.2 | 0.98 | 0.97 | 1 | 0.94 | 1 | 1 | 1 | 0.9 | 1.01 | 1.01 | 0.99 | [12 empty] | 1
| 1 | 2 | 1.11 | 1.01 | 0.8086 | 0.89 | 0.94 | 1.11 | 0.98 | 1.2 | 0.97 | 1 | 1.02 | 0.99 | 0.98 | 1 | 1 | 0.89 | 1.03 | 1 | 1.01 | [12 empty] | 27
| 1 | 3 | 1.12 | 1.01 | 0.79886 | 0.83 | 0.94 | 1.09 | 0.95 | 1.2 | 0.92 | 0.98 | 1.01 | 0.96 | 0.97 | 1 | 0.99 | 0.92 | 1.05 | 1.01 | 1.01 | [12 empty] | 16
| 1 | 4 | 1.13 | 1.02 | 0.79733 | 0.76 | 0.94 | 1.05 | 1.01 | 1.2 | 0.96 | 0.98 | 0.98 | 1 | 0.95 | 1 | 0.99 | 0.93 | 1.04 | 1 | 1.05 | [12 empty] | 0.99546
| 1 | 5 | 1.07 | 0.98 | 0.77785 | 0.87 | 0.91 | 1.06 | 0.95 | 1.2 | 0.96 | 1.04 | 0.99 | 1.01 | 1 | 1 | 1.02 | 0.83 | 1.05 | 1.02 | 1.03 | [12 empty] | 87
| 1 | 6 | 1.1 | 1 | 0.78382 | 0.87 | 0.93 | 1.02 | 0.97 | 1.1 | 0.98 | 1 | 0.99 | 0.99 | 0.97 | 1 | 1 | 0.87 | 1.05 | 0.98 | 1.04
| Mean |  | 1.1083 | 1 | 0.79608 | 0.84333 | 0.93333 | 1.0583 | 0.97667 | 1.1833 | 0.96167 | 0.995 | 0.99833 | 0.98167 | 0.97833 | 1 | 1 | 0.89 | 1.0383 | 1.0033 | 1.0217
| Median |  | 1.115 | 1.005 | 0.79809 | 0.855 | 0.94 | 1.055 | 0.975 | 1.2 | 0.965 | 0.99 | 0.995 | 0.99 | 0.975 | 1 | 1 | 0.895 | 1.045 | 1.005 | 1.02
| Std Dev. |  | 0.02137 | 0.016733 | 0.01299 | 0.046332 | 0.012111 | 0.03656 | 0.025033 | 0.040825 | 0.022286 | 0.0251 | 0.01472 | 0.026394 | 0.019408 | 0 | 0.010954 | 0.036332 | 0.016021 | 0.013663 | 0.022286
| Rel.Std.Dev. |  | 1.928% | 1.673% | 1.632% | 5.494% | 1.298% | 3.455% | 2.563% | 3.45% | 2.317% | 2.523% | 1.474% | 2.689% | 1.984% | 0% | 1.095% | 4.082% | 1.543% | 1.362% | 2.181%
| PDM3 |  | 11.34% | 0.4559% | -20.03% | -15.28% | -6.241% | 6.316% | -1.888% | 18.87% | -3.395% | -0.04636% | 0.2885% | -1.386% | -1.721% | 0.4559% | 0.4559% | -10.59% | 4.307% | 0.7908% | 2.632%
| Z-Score (Absolute) |  | 2.87 | 0.13 | 5.03 | 3.84 | 1.56 | 1.6 | 0.46 | NA | 0.84 | 0 | 0.08 | 0.34 | 0.42 | NA | 0.13 | 2.66 | 1.1 | 0.21 | 0.67
Results from laboratories 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 240b (Certified Value 162 ppm) | [32 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 181.6 | 164.52 | 156.65 | 180.1 | 168.3049 | 160 | 169 | 155 | 165.3 | 158 | 156.4 | 169 | 161.5 | 163.5 | 170.565 | 156 | 176 | 162.6273 | 164 | 155.52 | 158 | 163.07 | 155.2776 | 156 | 168.9796 | 165 | 142.685 | 168 | 161.5 | 158.7 | 162 | 1
| 1 | 2 | 182.9 | 168.87 | 155.11 | 174.84 | 167.5892 | 166 | 169 | 155 | 164.8 | 159 | 153.5 | 165.5 | 160.5 | 168.5 | 164.851 | 161.5 | 168 | 166.0014 | 163 | 153.2 | 162 | 159.42 | 160.9212 | 156 | 166.5464 | 161 | 142.057 | 166 | 164.4 | 155.7 | 160 | 28
| 1 | 3 | 187.1 | 167.94 | 155.54 | 178.59 | 168.6348 | 161 | 170 | 158 | 161.6 | 159 | 155.7 | 170.5 | 161 | 169 | 160.718 | 159.5 | 177 | 168.1251 | 165 | 153.23 | 158 | 162.38 | 161.6304 | 159 | 169.7675 | 164.5 | 141.369 | 166 | 169.4 | 160.7 | 164 | 16
| 1 | 4 | 180.9 | 166.6 | 155.95 | 178.89 | 168.0865 | 156 | 167 | 160 | 168.4 | 160 | 155.5 | 167.5 | 164.5 | 163 | 162.506 | 163.5 | 173 | 162.1258 | 168 | 133.35 | 159 | 160.6 | 156.0504 | 157 | 161.3846 | 165.5 | 137.437 | 165 | 162.1 | 156.9 | 162 | 162.498
| 1 | 5 | 181.1 | 168.87 | 156.55 | 176.03 | 165.3815 | 159 | 168 | 156 | 162.3 | 160 | 163.1 | 169.5 | 165.5 | 166 | 160.357 | 161 | 170 | 161.2587 | 167 | 158.49 | 161 | 160.88 | 163.1896 | 154 | 171.4348 | 165.5 | 131.04 | 171 | 164.9 | 156.6 | 161 | 88
| 1 | 6 | 183.7 | 164.15 | 154.87 | 180.49 | 169.5831 | 166 | 166 | 156 | 161.2 | 157 | 156 | 169.5 | 166 | 165 | 162.664 | 163.5 | 171 | 160.1477 | 162 | 151.33 | 159 | 164.04 | 157.872 | 154 | 164.782 | 163 | 152.514 | 167 | 163.5 | 160 | 164
| Mean |  | 182.8833 | 166.825 | 155.7783 | 178.1567 | 167.93 | 161.3333 | 168.1667 | 156.6667 | 163.9333 | 158.8333 | 156.7 | 168.5833 | 163.1667 | 165.8333 | 163.6102 | 160.8333 | 172.5 | 163.381 | 164.8333 | 150.8533 | 159.5 | 161.7317 | 159.1569 | 156 | 167.1491 | 164.0833 | 141.1837 | 167.1667 | 164.3 | 158.1 | 162.1667
| Median |  | 182.25 | 167.27 | 155.745 | 178.74 | 168.1957 | 160.5 | 168.5 | 156 | 163.55 | 159 | 155.85 | 169.25 | 163 | 165.5 | 162.585 | 161.25 | 172 | 162.3766 | 164.5 | 153.215 | 159 | 161.63 | 159.3966 | 156 | 167.763 | 164.75 | 141.713 | 166.5 | 163.95 | 157.8 | 162
| Std Dev. |  | 2.3327 | 2.1036 | 0.73665 | 2.257 | 1.4146 | 3.9833 | 1.472 | 1.9664 | 2.7624 | 1.169 | 3.293 | 1.8005 | 2.4427 | 2.5033 | 3.7675 | 2.8225 | 3.5071 | 3.0488 | 2.3166 | 8.9202 | 1.6432 | 1.7257 | 3.2201 | 1.8974 | 3.6816 | 1.7725 | 7.0478 | 2.137 | 2.8178 | 2.009 | 1.6021
| Rel.Std.Dev. |  | 1.276% | 1.261% | 0.4729% | 1.267% | 0.8424% | 2.469% | 0.8753% | 1.255% | 1.685% | 0.736% | 2.101% | 1.068% | 1.497% | 1.51% | 2.303% | 1.755% | 2.033% | 1.866% | 1.405% | 5.913% | 1.03% | 1.067% | 2.023% | 1.216% | 2.203% | 1.08% | 4.992% | 1.278% | 1.715% | 1.271% | 0.9879%
| PDM3 |  | 12.54% | 2.663% | -4.135% | 9.636% | 3.343% | -0.7168% | 3.488% | -3.589% | 0.8833% | -2.255% | -3.568% | 3.745% | 0.4115% | 2.053% | 0.6844% | -1.024% | 6.155% | 0.5434% | 1.437% | -7.166% | -1.845% | -0.4716% | -2.056% | -3.999% | 2.862% | 0.9756% | -13.12% | 2.873% | 1.109% | -2.707% | -0.2039%
| Z-Score (Absolute) |  | 3.39 | 0.6 | 1.32 | 2.57 | 0.79 | 0.36 | 0.83 | 1.17 | 0.1 | 0.79 | 1.16 | 0.9 | 0.04 | 0.43 | 0.04 | 0.44 | 1.58 | 0 | 0.25 | 2.18 | 0.67 | 0.29 | 0.73 | 1.28 | 0.65 | 0.12 | 3.86 | 0.66 | 0.16 | 0.92 | 0.21
| Analytical results for Dy in OREAS 240b (Certified Value 1.93 ppm) | [32 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 2.023 | 1.821 | 1.651 | 2.2 | 2.39 | 1.7 | [25 empty] | 1
| 1 | 2 | 2.015 | 1.782 | 1.6412 | 2.1 | 2.5 | 1.6 | [25 empty] | 29
| 1 | 3 | 2.01 | 1.881 | 1.6437 | 2.1 | 2.31 | 1.7 | [25 empty] | 16
| 1 | 4 | 2.043 | 1.806 | 1.6133 | 2.1 | 2.34 | 1.7 | [25 empty] | 1.9331
| 1 | 5 | 2.039 | 1.875 | 1.575 | 2.2 | 2.32 | 1.5 | [25 empty] | 89
| 1 | 6 | 1.976 | 1.948 | 1.6086 | 2.2 | 2.34 | 1.5
| Mean |  | 2.0177 | 1.8522 | 1.6221 | 2.15 | 2.3667 | 1.6167
| Median |  | 2.019 | 1.848 | 1.6273 | 2.15 | 2.34 | 1.65
| Std Dev. |  | 0.024196 | 0.060931 | 0.028795 | 0.054772 | 0.070899 | 0.098319
| Rel.Std.Dev. |  | 1.199% | 3.29% | 1.775% | 2.548% | 2.996% | 6.082%
| PDM3 |  | 4.374% | -4.187% | -16.09% | 11.22% | 22.43% | -16.37%
| Z-Score (Absolute) |  | 0.21 | 0.21 | 0.8 | 0.55 | 1.1 | 0.81
| Analytical results for Er in OREAS 240b (Certified Value 1.11 ppm) | [32 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.178 | 1.031 | 1.0065 | 1.2 | 1.35 | 0.9 | [25 empty] | 1
| 1 | 2 | 1.137 | 1.03 | 0.99265 | 1.4 | 1.42 | 0.9 | [25 empty] | 30
| 1 | 3 | 1.161 | 1.056 | 0.9683 | 1.2 | 1.3 | 0.9 | [25 empty] | 16
| 1 | 4 | 1.163 | 1.013 | 0.98276 | 1.3 | 1.29 | 0.9 | [25 empty] | 1.1141
| 1 | 5 | 1.172 | 1.06 | 0.97071 | 1.3 | 1.31 | 0.85 | [25 empty] | 90
| 1 | 6 | 1.141 | 1.08 | 0.98391 | 1.3 | 1.27 | 0.8
| Mean |  | 1.1587 | 1.045 | 0.98415 | 1.2833 | 1.3233 | 0.875
| Median |  | 1.162 | 1.0435 | 0.98334 | 1.3 | 1.305 | 0.9
| Std Dev. |  | 0.016476 | 0.02456 | 0.01419 | 0.075277 | 0.054283 | 0.041833
| Rel.Std.Dev. |  | 1.422% | 2.35% | 1.442% | 5.866% | 4.102% | 4.781%
| PDM3 |  | 4.002% | -6.201% | -11.66% | 15.19% | 18.78% | -21.46%
| Z-Score (Absolute) |  | 0.26 | 0.26 | 0.53 | 0.82 | 1 | 1.02
| Analytical results for Eu in OREAS 240b (Certified Value 0.45 ppm) | [32 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 0.467 | 0.381 | 0.48322 | 0.5 | 0.53 | 0.4 | [25 empty] | 1
| 1 | 2 | 0.473 | 0.37 | 0.49411 | 0.5 | 0.58 | 0.3 | [25 empty] | 31
| 1 | 3 | 0.472 | 0.393 | 0.49267 | 0.5 | 0.53 | 0.3 | [25 empty] | 16
| 1 | 4 | 0.468 | 0.373 | 0.48223 | 0.5 | 0.53 | 0.4 | [25 empty] | 0.45028
| 1 | 5 | 0.475 | 0.395 | 0.46559 | 0.5 | 0.55 | 0.3 | [25 empty] | 91
| 1 | 6 | 0.467 | 0.407 | 0.47127 | 0.5 | 0.51 | 0.3
| Mean |  | 0.47033 | 0.3865 | 0.48151 | 0.5 | 0.53833 | 0.33333
| Median |  | 0.47 | 0.387 | 0.48272 | 0.5 | 0.53 | 0.3
| Std Dev. |  | 0.0034448 | 0.014279 | 0.011359 | 0 | 0.024014 | 0.05164
| Rel.Std.Dev. |  | 0.7324% | 3.695% | 2.359% | 0% | 4.461% | 15.49%
| PDM3 |  | 4.453% | -14.16% | 6.937% | 11.04% | 19.56% | -25.97%
| Z-Score (Absolute) |  | 0.09 | 1.39 | 0.09 | 0.38 | 0.97 | 2.22
| Analytical results for Fe in OREAS 240b (Certified Value 5.4 wt.%) | [32 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 5.57 | 5.356 | 3.8631 | 5.1889 | 5.5815 | 5.59 | 5.498 | 5.81 | 5.42 | 5.41 | 5.17 | 5.33 | 5.27 | 5.35 | 5.6009 | 5.3 | 5.42 | > 5 | 5.99 | 4.86 | 5.33 | 6.5 | 5.5672 | 4.8265 | 5.23 | 5.27 | 5.41 | 5.42 | 5.27 |  |  | 1
| 1 | 2 | 5.65 | 5.401 | 3.8177 | 5.1993 | 5.4604 | 5.78 | 5.638 | 5.78 | 5.48 | 5.36 | 5.28 | 5.21 | 5.2 | 5.28 | 5.4081 | 5.42 | 5.2 | > 5 | 5.74 | 4.91 | 5.42 | 6.64 | 5.8318 | 4.7445 | 5.24 | 5.37 | 5.58 | 5.33 | 5.14 |  |  | #N/A
| 1 | 3 | 5.62 | 5.435 | 3.831 | 5.1413 | 5.5032 | 5.52 | 5.645 | 5.94 | 5.21 | 5.37 | 5.17 | 5.32 | 5.18 | 5.32 | 5.3306 | 5.44 | 5.5 | > 5 | 5.74 | 4.71 | 5.32 | 7.21 | 5.6821 | 4.897 | 5.26 | 5.3 | 5.45 | 5.49 | 5.23 |  |  | 16
| 1 | 4 | 5.63 | 5.481 | 3.8053 | 5.1682 | 5.4988 | 5.39 | 5.547 | 5.98 | 5.41 | 5.38 | 5.21 | 5.24 | 5.33 | 5.3 | 5.3818 | 5.45 | 5.49 | > 5 | 5.9 | 4.23 | 5.36 | 5.49 | 5.4739 | 4.815 | 5.28 | 5.3 | 5.44 | 5.31 | 5.27 |  |  | 5.3988
| 1 | 5 | 5.49 | 5.417 | 3.8131 | 5.196 | 5.4041 | 5.73 | 5.575 | 5.84 | 5.31 | 5.35 | 5.18 | 5.32 | 5.32 | 5.32 | 5.3499 | 5.47 | 5.45 | > 5 | 5.63 | 4.95 | 5.4 | 5.66 | 5.6867 | 4.6875 | 5.26 | 5.3 | 5.41 | 5.4 | 5.2 |  |  | 92
| 1 | 6 | 5.57 | 5.316 | 3.79 | 5.157 | 5.5366 | 5.76 | 5.624 | 5.81 | 5.31 | 5.33 | 5.18 | 5.29 | 5.34 | 5.33 | 5.3532 | 5.4 | 5.38 | > 5 | 5.51 | 4.8 | 5.35 | 6.39 | 5.458 | 4.701 | 5.27 | 5.26 | 5.45 | 5.4 | 5.28
| Mean |  | 5.5883 | 5.401 | 3.82 | 5.1751 | 5.4974 | 5.6283 | 5.5878 | 5.86 | 5.3567 | 5.3667 | 5.1983 | 5.285 | 5.2733 | 5.3167 | 5.4041 | 5.4133 | 5.4067 |  | 5.7517 | 4.7433 | 5.3633 | 6.315 | 5.6166 | 4.7786 | 5.2567 | 5.3 | 5.4567 | 5.3917 | 5.2317
| Median |  | 5.595 | 5.409 | 3.8154 | 5.1785 | 5.501 | 5.66 | 5.5995 | 5.825 | 5.36 | 5.365 | 5.18 | 5.305 | 5.295 | 5.32 | 5.3675 | 5.43 | 5.435 |  | 5.74 | 4.83 | 5.355 | 6.445 | 5.6246 | 4.7797 | 5.26 | 5.3 | 5.445 | 5.4 | 5.25
| Std Dev. |  | 0.058109 | 0.058416 | 0.025092 | 0.023369 | 0.061187 | 0.15536 | 0.058301 | 0.080747 | 0.097912 | 0.027325 | 0.042622 | 0.049295 | 0.069186 | 0.024221 | 0.10017 | 0.060553 | 0.11057 |  | 0.17452 | 0.2653 | 0.039328 | 0.64127 | 0.14383 | 0.08139 | 0.018619 | 0.038471 | 0.06314 | 0.064936 | 0.054191
| Rel.Std.Dev. |  | 1.04% | 1.082% | 0.6568% | 0.4516% | 1.113% | 2.76% | 1.043% | 1.378% | 1.828% | 0.5092% | 0.8199% | 0.9327% | 1.312% | 0.4556% | 1.854% | 1.119% | 2.045% |  | 3.034% | 5.593% | 0.7333% | 10.15% | 2.561% | 1.703% | 0.3542% | 0.7259% | 1.157% | 1.204% | 1.036%
| PDM3 |  | 3.511% | 0.04112% | -29.24% | -4.143% | 1.827% | 4.252% | 3.502% | 8.543% | -0.7801% | -0.5948% | -3.713% | -2.108% | -2.324% | -1.521% | 0.09791% | 0.2696% | 0.1461% |  | 6.536% | -12.14% | -0.6566% | 16.97% | 4.035% | -11.49% | -2.632% | -1.83% | 1.072% | -0.1318% | -3.095%
| Z-Score (Absolute) |  | 1.17 | 0.12 | 8.73 | 1.14 | 0.66 | 1.4 | 1.17 | 2.69 | 0.13 | 0.07 | 1.01 | 0.53 | 0.59 | 0.35 | 0.14 | 0.19 | 0.15 | NA | 2.09 | 3.56 | 0.09 | 5.24 | 1.33 | 3.36 | 0.69 | 0.44 | 0.43 | 0.07 | 0.83
| Analytical results for Ga in OREAS 240b (Certified Value 9.33 ppm) | [32 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 9 | 9.9 | 8.4007 | 8.1929 | 8.93 | 10.7 | 9.8 | 9.4 | 10 | 9 | 9.41 | 9.3 | 9.07 | 8.78 | 6 | 10.1 | 8.4 | 10.6 | 11.8 | 9 | 9.08 | 11 | 9.3 | 7.31 | 8.36 | [6 empty] | 1
| 1 | 2 | 8.9 | 9.7 | 8.4671 | 7.9132 | 9.28 | 10.7 | 9.5 | 9.34 | 10 | 9 | 9.29 | 9.36 | 9 | 9.31 | 6 | 10 | 8.4 | 10.9 | 11.69 | 9 | 9.01 | 11.4 | 9.2 | 7.45 | 8.26 | [6 empty] | 32
| 1 | 3 | 8.8 | 9.8 | 8.4405 | 8.2242 | 8.95 | 10.5 | 9.7 | 9.05 | 10 | 9 | 9.27 | 9.24 | 9 | 9.27 | 6 | 9.7 | 8.1 | 10.6 | 12.04 | 9 | 9.09 | 11.3 | 9.8 | 7.45 | 8.33 | [6 empty] | 16
| 1 | 4 | 8.9 | 9.9 | 8.3308 | 8.1015 | 8.81 | 10.2 | 9.9 | 9.15 | 10 | 9 | 9.74 | 9.29 | 8.82 | 9.62 | 6 | 9.8 | 7.4 | 11.1 | 12.38 | 9 | 9.13 | 11.2 | 9.5 | 7.22 | 8.34 | [6 empty] | 9.3294
| 1 | 5 | 8.6 | 9.8 | 8.4096 | 7.7603 | 8.99 | 10.1 | 9.6 | 8.86 | 10 | 10 | 9.3 | 9.71 | 8.92 | 9.45 | 6 | 9.3 | 8.6 | 10.7 | 11.59 | 8 | 9.28 | 11.1 | 9.6 | 7.37 | 8.21 | [6 empty] | 93
| 1 | 6 | 8.7 | 9.8 | 8.435 | 8.2002 | 9.32 | 9.78 | 9.5 | 9.01 | 9 | 9 | 9.45 | 9.42 | 8.99 | 9.46 | 6 | 9.5 | 8.2 | 10.8 | 12.09 | 8 | 9.28 | 10.8 | 9.4 | 7.28 | 8.37
| Mean |  | 8.8167 | 9.8167 | 8.414 | 8.0654 | 9.0467 | 10.33 | 9.6667 | 9.135 | 9.8333 | 9.1667 | 9.41 | 9.3867 | 8.9667 | 9.315 | 6 | 9.7333 | 8.1833 | 10.7833 | 11.9317 | 8.6667 | 9.145 | 11.1333 | 9.4667 | 7.3467 | 8.3117
| Median |  | 8.85 | 9.8 | 8.4223 | 8.1472 | 8.97 | 10.35 | 9.65 | 9.1 | 10 | 9 | 9.355 | 9.33 | 8.995 | 9.38 | 6 | 9.75 | 8.3 | 10.75 | 11.92 | 9 | 9.11 | 11.15 | 9.45 | 7.34 | 8.335
| Std Dev. |  | 0.1472 | 0.075277 | 0.047064 | 0.18826 | 0.20559 | 0.36742 | 0.1633 | 0.2054 | 0.40825 | 0.40825 | 0.17697 | 0.17014 | 0.086178 | 0.29002 | 0 | 0.30111 | 0.42151 | 0.19408 | 0.29335 | 0.5164 | 0.11149 | 0.21602 | 0.21602 | 0.093524 | 0.063061
| Rel.Std.Dev. |  | 1.67% | 0.7668% | 0.5594% | 2.334% | 2.273% | 3.557% | 1.689% | 2.249% | 4.152% | 4.454% | 1.881% | 1.813% | 0.9611% | 3.113% | 0% | 3.094% | 5.151% | 1.8% | 2.459% | 5.958% | 1.219% | 1.94% | 2.282% | 1.273% | 0.7587%
| PDM3 |  | -5.496% | 5.223% | -9.813% | -13.55% | -3.031% | 10.72% | 3.615% | -2.084% | 5.401% | -1.745% | 0.8637% | 0.6136% | -3.888% | -0.1546% | -35.69% | 4.329% | -12.28% | 15.58% | 27.89% | -7.104% | -1.977% | 19.34% | 1.471% | -21.25% | -10.91%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 1.22 | 1.69 | 0.36 | 1.37 | 0.48 | 0.24 | NA | NA | 0.13 | 0.1 | 0.47 | 0 | NA | 0.57 | 1.53 | 1.99 | 3.54 | NA | 0.23 | 2.46 | 0.21 | 2.66 | 1.36
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 240b (Certified Value 1.81 ppm) | [32 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.932 | 1.488 | 1.9638 | 2 | 1.89 | 1.6 | [25 empty] | 1
| 1 | 2 | 1.918 | 1.463 | 1.9403 | 2 | 1.95 | 1.6 | [25 empty] | 33
| 1 | 3 | 1.922 | 1.531 | 1.9263 | 2.1 | 1.89 | 1.6 | [25 empty] | 16
| 1 | 4 | 1.954 | 1.461 | 1.8997 | 1.9 | 1.83 | 1.7 | [25 empty] | 1.8069
| 1 | 5 | 1.952 | 1.555 | 1.8917 | 2 | 1.88 | 1.5 | [25 empty] | 94
| 1 | 6 | 1.936 | 1.604 | 1.9321 | 2 | 1.84 | 1.5
| Mean |  | 1.9357 | 1.517 | 1.9256 | 2 | 1.88 | 1.5833
| Median |  | 1.934 | 1.5095 | 1.9292 | 2 | 1.885 | 1.6
| Std Dev. |  | 0.014935 | 0.056749 | 0.026607 | 0.063246 | 0.042895 | 0.075277
| Rel.Std.Dev. |  | 0.7716% | 3.741% | 1.382% | 3.162% | 2.282% | 4.754%
| PDM3 |  | 7.124% | -16.05% | 6.569% | 10.68% | 4.043% | -12.37%
| Z-Score (Absolute) |  | 0.34 | 4 | 0.24 | 1.01 | 0.24 | 3.31
| Analytical results for Ge in OREAS 240b (Certified Value 0.094 ppm) | [32 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.1 | 0.06 | 0.19115 | < 0.2 | < 0.1 | 0.07 | 2 | 0.1 | 0.05 | 0.08 | 0.1 | < 10 | < 0.1 | 0.12 | 0.12 | 0.2 | 0.09 | 0.1 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.07 | 0.19784 | < 0.2 | < 0.1 | 0.09 | 2 | 0.12 | 0.05 | 0.1 | 0.11 | < 10 | < 0.1 | 0.13 | 0.17 | 0.2 | 0.09 | 0.1 | [13 empty] | 34
| 1 | 3 | < 0.1 | 0.07 | 0.18959 | < 0.2 | < 0.1 | 0.08 | 2 | 0.11 | 0.05 | 0.08 | 0.13 | < 10 | < 0.1 | 0.12 | 0.17 | 0.2 | 0.09 | 0.1 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.07 | 0.18688 | < 0.2 | < 0.1 | 0.05 | 1.9 | 0.12 | 0.05 | 0.08 | 0.11 | < 10 | < 0.1 | 0.12 | 0.09 | 0.2 | 0.09 | 0.11 | [13 empty] | 0.094167
| 1 | 5 | < 0.1 | 0.06 | 0.17812 | < 0.2 | < 0.1 | 0.08 | 2 | 0.1 | 0.05 | 0.08 | 0.11 | < 10 | < 0.1 | 0.13 | 0.12 | 0.2 | 0.09 | 0.1 | [13 empty] | 95
| 1 | 6 | < 0.1 | 0.07 | 0.18895 | < 0.2 | < 0.1 | 0.04 | 1.9 | 0.11 | 0.05 | 0.08 | 0.12 | < 10 | < 0.1 | 0.11 | 0.15 | 0.2 | 0.09 | 0.1
| Mean |  |  | 0.066667 | 0.18875 |  |  | 0.068333 | 1.9667 | 0.11 | 0.05 | 0.083333 | 0.11333 |  |  | 0.12167 | 0.13667 | 0.2 | 0.09 | 0.10167
| Median |  |  | 0.07 | 0.18927 |  |  | 0.075 | 2 | 0.11 | 0.05 | 0.08 | 0.11 |  |  | 0.12 | 0.135 | 0.2 | 0.09 | 0.1
| Std Dev. |  |  | 0.005164 | 0.0064135 |  |  | 0.019408 | 0.05164 | 0.0089443 | 7.6012e-18 | 0.008165 | 0.010328 |  |  | 0.0075277 | 0.032042 | 3.0405e-17 | 1.5202e-17 | 0.0040825
| Rel.Std.Dev. |  |  | 7.746% | 3.398% |  |  | 28.4% | 2.626% | 8.131% | 1.52e-14% | 9.798% | 9.113% |  |  | 6.187% | 23.45% | 1.52e-14% | 1.689e-14% | 4.016%
| PDM3 |  |  | -29.2% | 100.4% |  |  | -27.43% | 1988% | 16.81% | -46.9% | -11.5% | 20.35% |  |  | 29.2% | 45.13% | 112.4% | -4.425% | 7.965%
| Z-Score (Absolute) |  | 1.1 | 0.69 | 2.3 | 0.12 | 1.1 | 0.65 | NA | 0.37 | 1.1 | 0.29 | 0.45 | 120.27 | 1.1 | 0.65 | 1.02 | NA | 0.12 | 0.16
Results from laboratories 1.13 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 240b (Certified Value 0.43 ppm) | [32 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 2
| 1 | 1 | 0.37 | 0.46 | 0.32022 | 0.4 | 0.32 | 0.41 | 0.56 | 0.3 | 0.4 | 0.44 | 0.43 | 0.4 | 0.47 | < 0.5 | 0.46 | 0.93 | 0.43 | 0.5 | 0.39 | [12 empty] | 1
| 1 | 2 | 0.34 | 0.46 | 0.34353 | 0.4 | 0.3 | 0.44 | 0.55 | 0.3 | 0.45 | 0.42 | 0.45 | 0.41 | 0.46 | < 0.5 | 0.45 | 0.89 | 0.44 | 0.49 | 0.41 | [12 empty] | 18
| 1 | 3 | 0.32 | 0.46 | 0.32418 | 0.4 | 0.28 | 0.45 | 0.56 | 0.4 | 0.43 | 0.43 | 0.43 | 0.4 | 0.44 | < 0.5 | 0.47 | 0.87 | 0.45 | 0.52 | 0.37 | [12 empty] | 16
| 1 | 4 | 0.36 | 0.48 | 0.32636 | 0.3 | 0.32 | 0.39 | 0.55 | 0.3 | 0.44 | 0.45 | 0.46 | 0.41 | 0.52 | < 0.5 | 0.45 | 0.75 | 0.44 | 0.5 | 0.39 | [12 empty] | 0.43377
| 1 | 5 | 0.35 | 0.46 | 0.32924 | 0.4 | 0.3 | 0.42 | 0.53 | 0.3 | 0.44 | 0.49 | 0.46 | 0.43 | 0.52 | < 0.5 | 0.47 | 0.69 | 0.42 | 0.52 | 0.37 | [12 empty] | 96
| 1 | 6 | 0.32 | 0.44 | 0.31897 | 0.4 | 0.3 | 0.37 | 0.54 | 0.3 | 0.41 | 0.43 | 0.43 | 0.42 | 0.46 | < 0.5 | 0.47 | 0.67 | 0.44 | 0.51 | 0.37
| Mean |  | 0.34333 | 0.46 | 0.32708 | 0.38333 | 0.30333 | 0.41333 | 0.54833 | 0.31667 | 0.42833 | 0.44333 | 0.44333 | 0.41167 | 0.47833 |  | 0.46167 | 0.8 | 0.43667 | 0.50667 | 0.38333
| Median |  | 0.345 | 0.46 | 0.32527 | 0.4 | 0.3 | 0.415 | 0.55 | 0.3 | 0.435 | 0.435 | 0.44 | 0.41 | 0.465 |  | 0.465 | 0.81 | 0.44 | 0.505 | 0.38
| Std Dev. |  | 0.020656 | 0.012649 | 0.0089105 | 0.040825 | 0.015055 | 0.030111 | 0.01169 | 0.040825 | 0.019408 | 0.025033 | 0.015055 | 0.01169 | 0.033714 |  | 0.0098319 | 0.11082 | 0.010328 | 0.012111 | 0.01633
| Rel.Std.Dev. |  | 6.016% | 2.75% | 2.724% | 10.65% | 4.963% | 7.285% | 2.132% | 12.89% | 4.531% | 5.647% | 3.396% | 2.84% | 7.048% |  | 2.13% | 13.85% | 2.365% | 2.39% | 4.26%
| PDM3 |  | -20.85% | 6.047% | -24.6% | -11.63% | -30.07% | -4.712% | 26.41% | -27% | -1.254% | 2.204% | 2.204% | -5.096% | 10.27% |  | 6.431% | 84.43% | 0.6676% | 16.81% | -11.63%
| Z-Score (Absolute) |  | 1.51 | 0.38 | 1.77 | NA | 2.16 | 0.38 | 1.81 | NA | 0.13 | 0.11 | 0.11 | 0.4 | 0.67 | 3.02 | 0.4 | 5.88 | 0 | 1.13 | 0.86
Results from laboratories 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 240b (Indicative Value 0.062 ppm) | [32 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | < 0.01 | 0.02 | < 1 | 0.13 | < 1 | < 0.1 | < 5 | 0.686 | 0.04 | 0.03 | < 0.01 | 0.03 | < 1 | 0.12 | 0.006 | 0.09 | < 0.1 | 0.05 | < 0.08 | 0.05 | 0.191 | 0.38 | < 0.01 | 0.022 | 0.49 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.02 | < 1 | 0.1 | < 1 | < 0.1 | < 5 | 0.707 | 0.01 | 0.02 | < 0.01 | 0.04 | < 1 | 0.11 | 0.011 | 0.07 | < 0.1 | 0.05 | < 0.08 | 0.03 | 0.183 | 0.39 | < 0.01 | 0.022 | 0.5 | [6 empty] | 5
| 1 | 3 | < 0.01 | 0.02 | < 1 | 0.08 | < 1 | < 0.1 | < 5 | 0.656 | 0.04 | 0.02 | < 0.01 | 0.03 | < 1 | 0.12 | 0.013 | 0.08 | 0.11 | 0.05 | < 0.08 | 0.05 | 0.184 | 0.4 | < 0.01 | 0.022 | 0.49 | [6 empty] | 16
| 1 | 4 | < 0.01 | 0.02 | < 1 | 0.09 | < 1 | < 0.1 | < 5 | 0.644 | 0.03 | 0.06 | < 0.01 | 0.04 | < 1 | 0.13 | 0.007 | 0.07 | < 0.1 | 0.05 | < 0.08 | 0.04 | 0.162 | 0.4 | < 0.01 | 0.021 | 0.47 | [6 empty] | 0.062154
| 1 | 5 | < 0.01 | 0.02 | < 1 | 0.07 | < 1 | < 0.1 | < 5 | 0.703 | 0.05 | 0.02 | < 0.01 | 0.06 | < 1 | 0.13 | 0.009 | 0.08 | < 0.1 | 0.05 | < 0.08 | 0.04 | 0.184 | 0.39 | < 0.01 | 0.02 | 0.75 | [6 empty] | 11
| 1 | 6 | < 0.01 | 0.02 | < 1 | 0.04 | < 1 | < 0.1 | < 5 | 0.655 | 0.01 | 0.03 | < 0.01 | 0.03 | < 1 | 0.11 | 0.009 | 0.09 | < 0.1 | 0.05 | < 0.08 | 0.04 | 0.165 | 0.4 | < 0.01 | 0.023 | 0.66
| Mean |  |  | 0.02 |  | 0.085 |  |  |  | 0.67517 | 0.03 | 0.03 |  | 0.038333 |  | 0.12 | 0.0091667 | 0.08 | 0.11 | 0.05 |  | 0.041667 | 0.17817 | 0.39333 |  | 0.021667 | 0.56
| Median |  |  | 0.02 |  | 0.085 |  |  |  | 0.671 | 0.035 | 0.025 |  | 0.035 |  | 0.12 | 0.009 | 0.08 | 0.11 | 0.05 |  | 0.04 | 0.1835 | 0.395 |  | 0.022 | 0.495
| Std Dev. |  |  | 0 |  | 0.030166 |  |  |  | 0.027022 | 0.016733 | 0.015492 |  | 0.01169 |  | 0.0089443 | 0.0025626 | 0.0089443 |  | 7.6012e-18 |  | 0.0075277 | 0.011754 | 0.008165 |  | 0.0010328 | 0.11628
| Rel.Std.Dev. |  |  | 0% |  | 35.49% |  |  |  | 4.002% | 55.78% | 51.64% |  | 30.5% |  | 7.454% | 27.96% | 11.18% |  | 1.52e-14% |  | 18.07% | 6.597% | 2.076% |  | 4.767% | 20.76%
| PDM3 |  |  | -67.82% |  | 36.76% |  |  |  | 986.3% | -51.73% | -51.73% |  | -38.33% |  | 93.07% | -85.25% | 28.71% | 76.98% | -19.55% |  | -32.96% | 186.7% | 532.8% |  | -65.14% | 801%
| Z-Score (Absolute) |  | 0.74 | 0.5 | 7.43 | 0.58 | 7.43 | 0 | 40.46 | 10.32 | 0.33 | 0.33 | 0.74 | 0.19 | 7.43 | 1.16 | 0.67 | 0.5 | 0.17 | 0 | 0.17 | 0.14 | 2.12 | 5.67 | 0.74 | 0.47 | 8.42
| Analytical results for Ho in OREAS 240b (Certified Value 0.37 ppm) | [32 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 0.403 | 0.378 | 0.33976 | 0.4 | 0.51 | 0.3 | [25 empty] | 1
| 1 | 2 | 0.398 | 0.368 | 0.32992 | 0.4 | 0.52 | 0.35 | [25 empty] | 19
| 1 | 3 | 0.399 | 0.392 | 0.33261 | 0.4 | 0.48 | 0.35 | [25 empty] | 16
| 1 | 4 | 0.397 | 0.367 | 0.32984 | 0.4 | 0.48 | 0.4 | [25 empty] | 0.37267
| 1 | 5 | 0.399 | 0.384 | 0.33327 | 0.4 | 0.49 | 0.35 | [25 empty] | 97
| 1 | 6 | 0.398 | 0.396 | 0.33556 | 0.4 | 0.48 | 0.35
| Mean |  | 0.399 | 0.38083 | 0.33349 | 0.4 | 0.49333 | 0.35
| Median |  | 0.3985 | 0.381 | 0.33294 | 0.4 | 0.485 | 0.35
| Std Dev. |  | 0.0020976 | 0.012073 | 0.0037543 | 6.0809e-17 | 0.017512 | 0.031623
| Rel.Std.Dev. |  | 0.5257% | 3.17% | 1.126% | 1.52e-14% | 3.55% | 9.035%
| PDM3 |  | 7.067% | 2.192% | -10.51% | 7.335% | 32.38% | -6.082%
| Z-Score (Absolute) |  | 0.24 | 0.24 | 1.52 | 0.27 | 2.79 | 1.08
| Analytical results for In in OREAS 240b (Certified Value 0.038 ppm) | [32 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.04 | 0.04 | 0.033787 | 0.04 | 0.05 | 0.05 | < 5 | 0.04 | 0.036 | 0.036 | 0.035 | 0.035 | < 10 | < 0.05 | 0.64 | < 0.1 | 0.03973 | 0.039 | 0.04 | 0.04 | 0.036 | 0.033 | [9 empty] | 1
| 1 | 2 | 0.04 | 0.04 | 0.030209 | 0.04 | < 0.05 | 0.04 | < 5 | 0.04 | 0.035 | 0.038 | 0.041 | 0.038 | < 10 | < 0.05 | 0.48 | < 0.1 | 0.04048 | 0.037 | 0.043 | 0.04 | 0.036 | 0.032 | [9 empty] | 20
| 1 | 3 | 0.04 | 0.04 | 0.031634 | 0.04 | 0.05 | 0.04 | < 5 | 0.04 | 0.031 | 0.036 | 0.035 | 0.038 | < 10 | < 0.05 | 1.7 | < 0.1 | 0.0392 | 0.04 | 0.043 | 0.04 | 0.034 | 0.033 | [9 empty] | 16
| 1 | 4 | 0.04 | 0.04 | 0.030044 | 0.04 | < 0.05 | 0.05 | < 5 | 0.04 | 0.041 | 0.037 | 0.037 | 0.038 | < 10 | < 0.05 | 0.91 | < 0.1 | 0.03509 | 0.038 | 0.037 | 0.04 | 0.036 | 0.034 | [9 empty] | 0.038395
| 1 | 5 | 0.04 | 0.04 | 0.029945 | 0.06 | 0.05 | 0.04 | < 5 | 0.04 | 0.042 | 0.039 | 0.037 | 0.036 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.03713 | 0.04 | 0.045 | 0.04 | 0.036 | 0.031 | [9 empty] | 98
| 1 | 6 | 0.03 | 0.04 | 0.030507 | 0.04 | < 0.05 | 0.03 | < 5 | 0.04 | 0.04 | 0.038 | 0.036 | 0.034 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.04292 | 0.039 | 0.039 | 0.04 | 0.035 | 0.036
| Mean |  | 0.038333 | 0.04 | 0.031021 | 0.043333 | 0.05 | 0.041667 |  | 0.04 | 0.0375 | 0.037333 | 0.036833 | 0.0365 |  |  | 0.9325 |  | 0.039092 | 0.038833 | 0.041167 | 0.04 | 0.0355 | 0.033167
| Median |  | 0.04 | 0.04 | 0.030358 | 0.04 | 0.05 | 0.04 |  | 0.04 | 0.038 | 0.0375 | 0.0365 | 0.037 |  |  | 0.775 |  | 0.039465 | 0.039 | 0.0415 | 0.04 | 0.036 | 0.033
| Std Dev. |  | 0.0040825 | 0 | 0.0014876 | 0.008165 | 8.4984e-18 | 0.0075277 |  | 0 | 0.0042308 | 0.0012111 | 0.0022286 | 0.0017607 |  |  | 0.54156 |  | 0.0027153 | 0.001169 | 0.0029944 | 0 | 0.00083666 | 0.0017224
| Rel.Std.Dev. |  | 10.65% | 0% | 4.795% | 18.84% | 1.7e-14% | 18.07% |  | 0% | 11.28% | 3.244% | 6.051% | 4.824% |  |  | 58.08% |  | 6.946% | 3.01% | 7.274% | 0% | 2.357% | 5.193%
| PDM3 |  | -0.1606% | 4.18% | -19.21% | 12.86% | 30.23% | 8.521% |  | 4.18% | -2.331% | -2.765% | -4.067% | -4.936% |  |  | 2329% |  | 1.814% | 1.142% | 7.219% | 4.18% | -7.54% | -13.62%
| Z-Score (Absolute) |  | 0.18 | 0.3 | 2.29 | 1.26 | 0.42 | 0.78 | 711.21 | 0.3 | 0.42 | 0.47 | 0.62 | 0.71 | 1433.69 | 4.04 | 173.21 | 3.19 | 0.04 | 0.04 | 0.64 | 0.3 | 1 | 1.67
| Analytical results for Ir in OREAS 240b (Indicative Value < 0.003 ppm) | [32 empty] | Indicative
Ir | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.27 | 1.28 | [29 empty] | ppm
|  |  | AR*MS | AR*MS | [29 empty] | 3
|  |  | 25g | 0.5g | [29 empty] | 3
| 1 | 1 | N.A. | < 0.003 | [29 empty] | 1
| 1 | 2 | N.A. | < 0.003 | [29 empty] | 6
| 1 | 3 | N.A. | < 0.003 | [29 empty] | 16
| 1 | 4 | N.A. | < 0.003 | [29 empty] | < 0.003
| 1 | 5 | N.A. | < 0.003 | [29 empty] | 12
| 1 | 6 | N.A. | < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 240b (Certified Value 0.209 wt.%) | [32 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.22 | 0.191 | 0.25792 | 0.22018 | 0.20363 | 0.18 | 0.2142 | 0.238 | 0.2167 | 0.22 | 0.21 | 0.2 | 0.2 | 0.2 | 0.22133 | 0.2 | 0.19 | 0.22934 | 0.222 | 0.21 | 0.21 | 0.19875 | 0.19622 | 0.19 | 0.2 | 0.21 | 0.23 | 0.22 | 0.21 |  |  | 1
| 1 | 2 | 0.22 | 0.194 | 0.25449 | 0.22005 | 0.20022 | 0.18 | 0.2192 | 0.23 | 0.2231 | 0.22 | 0.22 | 0.2 | 0.2 | 0.2 | 0.21569 | 0.2 | 0.18 | 0.21991 | 0.211 | 0.21 | 0.21 | 0.19862 | 0.21301 | 0.1885 | 0.2 | 0.2 | 0.24 | 0.21 | 0.21 |  |  | #N/A
| 1 | 3 | 0.22 | 0.195 | 0.25313 | 0.21544 | 0.20266 | 0.18 | 0.22 | 0.234 | 0.2113 | 0.22 | 0.22 | 0.2 | 0.2 | 0.2 | 0.21369 | 0.2 | 0.19 | 0.22478 | 0.213 | 0.2 | 0.21 | 0.2023 | 0.20946 | 0.193 | 0.2 | 0.2 | 0.24 | 0.22 | 0.21 |  |  | 16
| 1 | 4 | 0.22 | 0.197 | 0.25542 | 0.21363 | 0.19923 | 0.17 | 0.2158 | 0.245 | 0.2192 | 0.22 | 0.23 | 0.2 | 0.2 | 0.2 | 0.21658 | 0.2 | 0.19 | 0.22969 | 0.221 | 0.18 | 0.21 | 0.20024 | 0.19908 | 0.191 | 0.2 | 0.2 | 0.23 | 0.21 | 0.21 |  |  | 0.20941
| 1 | 5 | 0.21 | 0.195 | 0.25832 | 0.21674 | 0.19895 | 0.18 | 0.2158 | 0.236 | 0.2131 | 0.22 | 0.22 | 0.2 | 0.21 | 0.2 | 0.21374 | 0.21 | 0.19 | 0.21306 | 0.22 | 0.22 | 0.21 | 0.19935 | 0.2033 | 0.185 | 0.2 | 0.21 | 0.23 | 0.22 | 0.21 |  |  | 99
| 1 | 6 | 0.22 | 0.192 | 0.25129 | 0.21624 | 0.20256 | 0.17 | 0.2192 | 0.239 | 0.2128 | 0.21 | 0.22 | 0.2 | 0.21 | 0.2 | 0.21416 | 0.21 | 0.18 | 0.21155 | 0.212 | 0.19 | 0.21 | 0.20323 | 0.19931 | 0.1855 | 0.2 | 0.2 | 0.23 | 0.22 | 0.21
| Mean |  | 0.21833 | 0.194 | 0.25509 | 0.21705 | 0.20121 | 0.17667 | 0.21737 | 0.237 | 0.21603 | 0.21833 | 0.22 | 0.2 | 0.20333 | 0.2 | 0.21587 | 0.20333 | 0.18667 | 0.22139 | 0.2165 | 0.20167 | 0.21 | 0.20042 | 0.2034 | 0.18883 | 0.2 | 0.20333 | 0.23333 | 0.21667 | 0.21
| Median |  | 0.22 | 0.1945 | 0.25496 | 0.21649 | 0.20139 | 0.18 | 0.2175 | 0.237 | 0.2149 | 0.22 | 0.22 | 0.2 | 0.2 | 0.2 | 0.21492 | 0.2 | 0.19 | 0.22235 | 0.2165 | 0.205 | 0.21 | 0.1998 | 0.20131 | 0.18925 | 0.2 | 0.2 | 0.23 | 0.22 | 0.21
| Std Dev. |  | 0.0040825 | 0.0021909 | 0.0027272 | 0.0026014 | 0.0019895 | 0.005164 | 0.0023914 | 0.0050596 | 0.0045138 | 0.0040825 | 0.0063246 | 3.0405e-17 | 0.005164 | 3.0405e-17 | 0.0029183 | 0.005164 | 0.005164 | 0.0078985 | 0.00501 | 0.01472 | 0 | 0.0019305 | 0.0065731 | 0.0031411 | 3.0405e-17 | 0.005164 | 0.005164 | 0.005164 | 0
| Rel.Std.Dev. |  | 1.87% | 1.129% | 1.069% | 1.199% | 0.9888% | 2.923% | 1.1% | 2.135% | 2.089% | 1.87% | 2.875% | 1.52e-14% | 2.54% | 1.52e-14% | 1.352% | 2.54% | 2.766% | 3.568% | 2.314% | 7.299% | 0% | 0.9633% | 3.232% | 1.663% | 1.52e-14% | 2.54% | 2.213% | 2.383% | 0%
| PDM3 |  | 4.262% | -7.358% | 21.82% | 3.647% | -3.917% | -15.64% | 3.8% | 13.18% | 3.163% | 4.262% | 5.057% | -4.493% | -2.901% | -4.493% | 3.083% | -2.901% | -10.86% | 5.72% | 3.386% | -3.697% | 0.2822% | -4.295% | -2.871% | -9.826% | -4.493% | -2.901% | 11.42% | 3.466% | 0.2822%
| Z-Score (Absolute) |  | 0.67 | 1.29 | 3.65 | 0.57 | 0.71 | 2.7 | 0.6 | 2.18 | 0.49 | 0.67 | 0.81 | 0.81 | 0.54 | 0.81 | 0.47 | 0.54 | 1.89 | 0.92 | 0.53 | 0.67 | 0 | 0.78 | 0.53 | 1.71 | 0.81 | 0.54 | 1.89 | 0.54 | 0
| Analytical results for La in OREAS 240b (Certified Value 5.49 ppm) | [32 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 5.1 | 5.479 | 4.658 | 5.81 | 5.7663 | 5.4 | 3 | 6.14 | 5.373 | 5 | 6 | 5.2 | 5.6 | 5.2 | 5 | 5 | 5.56 | 5.1 | 5.3 | 5.5 | 5.8 | 5.8 | 5.2 | 5.6 | 5.5 | [6 empty] | 1
| 1 | 2 | 5.2 | 5.43 | 4.667 | 5.8 | 5.7381 | 5.5 | 3 | 6.02 | 5.468 | 5 | 6 | 5.2 | 5.7 | 5.3 | 5.3 | 5 | 5.51 | 4.9 | 5.4 | 5.5 | 5.6 | 5.9 | 5.3 | 5.5 | 5.4 | [6 empty] | 21
| 1 | 3 | 5.2 | 5.435 | 4.744 | 5.87 | 5.7237 | 5.4 | 4 | 5.97 | 5.228 | 5 | 6 | 5 | 5.7 | 5.3 | 5.2 | 5 | 5.4 | 5 | 5.3 | 5.5 | 5.8 | 5.9 | 5.5 | 5.7 | 5.5 | [6 empty] | 16
| 1 | 4 | 5.1 | 5.565 | 4.646 | 5.89 | 5.743 | 5.2 | 3 | 6.17 | 5.381 | 5 | 6 | 5.3 | 5.7 | 5.1 | 5.4 | 5 | 5.62 | 4.5 | 5.4 | 5.5 | 5.8 | 5.9 | 5.4 | 5.5 | 5.6 | [6 empty] | 5.4853
| 1 | 5 | 5.1 | 5.469 | 4.845 | 5.76 | 5.7058 | 5.3 | 1 | 6 | 5.349 | 5 | 6 | 5.2 | 5.6 | 5.3 | 5.4 | 5 | 5.23 | 5.1 | 5.3 | 5.5 | 5.9 | 5.7 | 5.4 | 5.7 | 5.4 | [6 empty] | 100
| 1 | 6 | 5.1 | 5.402 | 5.107 | 5.84 | 5.7839 | 5.6 | 2 | 6.02 | 5.262 | 5 | 6 | 5.3 | 5.6 | 5.3 | 5.4 | 5 | 5.3 | 5 | 5.4 | 5.5 | 5.9 | 5.7 | 5.4 | 5.5 | 5.5
| Mean |  | 5.1333 | 5.4633 | 4.7778 | 5.8283 | 5.7435 | 5.4 | 2.6667 | 6.0533 | 5.3435 | 5 | 6 | 5.2 | 5.65 | 5.25 | 5.2833 | 5 | 5.4367 | 4.9333 | 5.35 | 5.5 | 5.8 | 5.8167 | 5.3667 | 5.5833 | 5.4833
| Median |  | 5.1 | 5.452 | 4.7055 | 5.825 | 5.7405 | 5.4 | 3 | 6.02 | 5.361 | 5 | 6 | 5.2 | 5.65 | 5.3 | 5.35 | 5 | 5.455 | 5 | 5.35 | 5.5 | 5.8 | 5.85 | 5.4 | 5.55 | 5.5
| Std Dev. |  | 0.05164 | 0.057078 | 0.17778 | 0.047924 | 0.028271 | 0.14142 | 1.0328 | 0.081404 | 0.086936 | 0 | 0 | 0.10954 | 0.054772 | 0.083666 | 0.16021 | 0 | 0.15293 | 0.22509 | 0.054772 | 0 | 0.10954 | 0.098319 | 0.10328 | 0.098319 | 0.075277
| Rel.Std.Dev. |  | 1.006% | 1.045% | 3.721% | 0.8223% | 0.4922% | 2.619% | 38.73% | 1.345% | 1.627% | 0% | 0% | 2.107% | 0.9694% | 1.594% | 3.032% | 0% | 2.813% | 4.563% | 1.024% | 0% | 1.889% | 1.69% | 1.924% | 1.761% | 1.373%
| PDM3 |  | -6.416% | -0.3998% | -12.9% | 6.254% | 4.707% | -1.554% | -51.38% | 10.36% | -2.584% | -8.847% | 9.384% | -5.201% | 3.003% | -4.289% | -3.681% | -8.847% | -0.886% | -10.06% | -2.466% | 0.2686% | 5.738% | 6.042% | -2.162% | 1.788% | -0.03522%
| Z-Score (Absolute) |  | 1.1 | 0.1 | 2.38 | 1.41 | 1.11 | 0.13 | NA | 2.23 | 0.34 | NA | NA | 0.85 | 0.77 | 0.67 | 0.55 | NA | 0 | 1.82 | 0.31 | 0.23 | 1.31 | 1.37 | 0.25 | 0.53 | 0.17
Results from laboratories 1.08, 1.12, 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 240b (Certified Value 10.6 ppm) | [32 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 1 | 10.23 | 17.74 | 3.9995 | 10.6422 | 10.2 | 8 | 12 | 10.6 | 11 | 13.5 | 11.3 | 10.9 | 11.4 | 10.5 | < 5 | 8.5041 | 10.7 | 10 | 9 | 11.1 | 10.4 | 12 | 10.5 | 9.9 | [6 empty] | 1
| 1 | 2 | 1 | 10.22 | 17.77 | 4.0017 | 10.4661 | 10.5 | 8 | 11.6 | 10.6 | 11 | 12.9 | 11 | 10.8 | 11.2 | 10.7 | < 5 | 9.0259 | 10.6 | 11 | 9 | 11.1 | 10 | 12 | 10.5 | 9.7 | [6 empty] | #N/A
| 1 | 3 | 2 | 10.2 | 17.09 | 4.0071 | 10.4794 | 10.2 | 8 | 11.9 | 10.5 | 11 | 13.1 | 11.3 | 10.7 | 11.3 | 10.9 | < 5 | 9.9393 | 10.3 | 10 | 10 | 11.2 | 10 | 12 | 10.7 | 9.9 | [6 empty] | 16
| 1 | 4 | 1 | 10.23 | 17.65 | 3.9729 | 10.3635 | 10.2 | 8 | 12.1 | 10.5 | 11 | 13.2 | 11.1 | 11 | 11.2 | 10.7 | < 5 | 9.7142 | 10.2 | 10 | 10 | 11.1 | 10.3 | 12 | 10.2 | 10.3 | [6 empty] | 10.5696
| 1 | 5 | 2 | 10.27 | 18.03 | 3.9931 | 10.4379 | 10.3 | 8 | 11.5 | 10.1 | 11 | 12.9 | 11.3 | 11.1 | 11.1 | 10.9 | < 5 | 8.9126 | 9.6 | 10 | 9 | 11.1 | 10.1 | 12 | 10.4 | 10.1 | [6 empty] | 101
| 1 | 6 | 2 | 10.1 | 17.5 | 3.9725 | 10.5485 | 10.6 | 8 | 11.4 | 10.3 | 11 | 13 | 11.2 | 11.1 | 11.2 | 10.8 | < 5 | 9.0576 | 9.5 | 11 | 9 | 11.1 | 10 | 12 | 10.4 | 10.3
| Mean |  | 1.5 | 10.2083 | 17.63 | 3.9911 | 10.4896 | 10.3333 | 8 | 11.75 | 10.4333 | 11 | 13.1 | 11.2 | 10.9333 | 11.2333 | 10.75 |  | 9.1923 | 10.15 | 10.3333 | 9.3333 | 11.1167 | 10.1333 | 12 | 10.45 | 10.0333
| Median |  | 1.5 | 10.225 | 17.695 | 3.9963 | 10.4727 | 10.25 | 8 | 11.75 | 10.5 | 11 | 13.05 | 11.25 | 10.95 | 11.2 | 10.75 |  | 9.0417 | 10.25 | 10 | 9 | 11.1 | 10.05 | 12 | 10.45 | 10
| Std Dev. |  | 0.54772 | 0.057764 | 0.31642 | 0.014968 | 0.09593 | 0.17512 | 0 | 0.2881 | 0.19664 | 0 | 0.22804 | 0.12649 | 0.1633 | 0.10328 | 0.15166 |  | 0.53445 | 0.501 | 0.5164 | 0.5164 | 0.040825 | 0.17512 | 0 | 0.16432 | 0.24221
| Rel.Std.Dev. |  | 36.51% | 0.5659% | 1.795% | 0.375% | 0.9145% | 1.695% | 0% | 2.452% | 1.885% | 0% | 1.741% | 1.129% | 1.494% | 0.9194% | 1.411% |  | 5.814% | 4.936% | 4.997% | 5.533% | 0.3672% | 1.728% | 0% | 1.572% | 2.414%
| PDM3 |  | -85.81% | -3.418% | 66.8% | -62.24% | -0.7572% | -2.236% | -24.31% | 11.17% | -1.29% | 4.072% | 23.94% | 5.964% | 3.441% | 6.279% | 1.706% |  | -13.03% | -3.97% | -2.236% | -11.7% | 5.175% | -4.128% | 13.53% | -1.132% | -5.074%
| Z-Score (Absolute) |  | NA | 0.34 | 10.02 | 9.01 | 0.06 | 0.16 | NA | 1.81 | 0.02 | NA | 3.7 | 1.05 | 0.67 | 1.09 | 0.42 | 11.09 | 1.75 | 0.42 | NA | NA | 0.93 | 0.44 | NA | 0 | 0.58
Results from laboratories 1.01, 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 240b (Certified Value 0.13 ppm) | [32 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | 1.32 | [24 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [24 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [24 empty] | 2
| 1 | 1 | 0.12 | 0.14 | 0.125 | 0.86327 | 0.1 | 0.1 | 0.12 | [24 empty] | 1
| 1 | 2 | 0.13 | 0.136 | 0.124 | 0.86467 | 0.2 | 0.1 | 0.12 | [24 empty] | 22
| 1 | 3 | 0.13 | 0.137 | 0.127 | 0.84232 | 0.2 | < 0.1 | 0.13 | [24 empty] | 16
| 1 | 4 | 0.12 | 0.142 | 0.126 | 0.83385 | 0.1 | 0.1 | 0.12 | [24 empty] | 0.12961
| 1 | 5 | 0.12 | 0.138 | 0.127 | 0.83351 | 0.2 | < 0.1 | 0.12 | [24 empty] | 102
| 1 | 6 | 0.12 | 0.138 | 0.134 | 0.82175 | 0.2 | < 0.1 | 0.13
| Mean |  | 0.12333 | 0.1385 | 0.12717 | 0.84323 | 0.16667 | 0.1 | 0.12333
| Median |  | 0.12 | 0.138 | 0.1265 | 0.83808 | 0.2 | 0.1 | 0.12
| Std Dev. |  | 0.005164 | 0.0021679 | 0.0035449 | 0.017357 | 0.05164 | 1.6997e-17 | 0.005164
| Rel.Std.Dev. |  | 4.187% | 1.565% | 2.788% | 2.058% | 30.98% | 1.7e-14% | 4.187%
| PDM3 |  | -4.839% | 6.863% | -1.882% | 550.6% | 28.6% | -22.84% | -4.839%
| Z-Score (Absolute) |  | 0.23 | 0.67 | 0 | 42.6 | 2.35 | 3.1 | 0.23
| Analytical results for Mg in OREAS 240b (Certified Value 1.7 wt.%) | [32 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.74 | 1.625 | 1.7269 | 1.7988 | 1.7295 | 1.79 | 1.8273 | 1.86 | 1.74 | 1.75 | 1.63 | 1.68 | 1.63 | 1.62 | 1.748 | 1.69 | 1.58 | 1.8399 | 1.82 | 1.63 | 1.69 | 2.2825 | 1.493 | 1.543 | 1.65 | 1.69 | 1.66 | 1.7 | 1.74 |  |  | 1
| 1 | 2 | 1.76 | 1.637 | 1.7199 | 1.7883 | 1.7075 | 1.85 | 1.8816 | 1.82 | 1.74 | 1.79 | 1.65 | 1.63 | 1.66 | 1.61 | 1.6827 | 1.74 | 1.52 | 1.784 | 1.76 | 1.64 | 1.72 | 2.3457 | 1.5745 | 1.5215 | 1.65 | 1.69 | 1.72 | 1.68 | 1.73 |  |  | #N/A
| 1 | 3 | 1.76 | 1.666 | 1.6876 | 1.7977 | 1.7022 | 1.73 | 1.8695 | 1.83 | 1.66 | 1.77 | 1.63 | 1.67 | 1.65 | 1.62 | 1.6656 | 1.74 | 1.6 | 1.8516 | 1.77 | 1.6 | 1.68 | 2.5559 | 1.547 | 1.557 | 1.65 | 1.68 | 1.68 | 1.72 | 1.79 |  |  | 16
| 1 | 4 | 1.77 | 1.652 | 1.7088 | 1.771 | 1.6817 | 1.72 | 1.8454 | 1.88 | 1.71 | 1.77 | 1.63 | 1.64 | 1.68 | 1.61 | 1.683 | 1.72 | 1.6 | 1.7505 | 1.81 | 1.43 | 1.69 | 1.9295 | 1.4571 | 1.547 | 1.67 | 1.67 | 1.67 | 1.67 | 1.76 |  |  | 1.696
| 1 | 5 | 1.72 | 1.644 | 1.7432 | 1.7727 | 1.7005 | 1.73 | 1.8514 | 1.84 | 1.71 | 1.78 | 1.62 | 1.66 | 1.68 | 1.61 | 1.6736 | 1.74 | 1.59 | 1.7666 | 1.75 | 1.69 | 1.71 | 1.9961 | 1.5194 | 1.5005 | 1.66 | 1.7 | 1.66 | 1.7 | 1.74 |  |  | 103
| 1 | 6 | 1.74 | 1.63 | 1.6986 | 1.797 | 1.7217 | 1.83 | 1.8635 | 1.84 | 1.7 | 1.72 | 1.62 | 1.65 | 1.65 | 1.62 | 1.673 | 1.72 | 1.58 | 1.7137 | 1.73 | 1.6 | 1.69 | 1.9295 | 1.4818 | 1.4965 | 1.66 | 1.67 | 1.67 | 1.7 | 1.75
| Mean |  | 1.7483 | 1.6423 | 1.7142 | 1.7876 | 1.7072 | 1.775 | 1.8564 | 1.845 | 1.71 | 1.7633 | 1.63 | 1.655 | 1.6583 | 1.615 | 1.6876 | 1.725 | 1.5783 | 1.7844 | 1.7733 | 1.5983 | 1.6967 | 2.1732 | 1.5121 | 1.5276 | 1.6567 | 1.6833 | 1.6767 | 1.695 | 1.7517
| Median |  | 1.75 | 1.6405 | 1.7144 | 1.7927 | 1.7048 | 1.76 | 1.8575 | 1.84 | 1.71 | 1.77 | 1.63 | 1.655 | 1.655 | 1.615 | 1.6781 | 1.73 | 1.585 | 1.7753 | 1.765 | 1.615 | 1.69 | 2.1393 | 1.5062 | 1.5322 | 1.655 | 1.685 | 1.67 | 1.7 | 1.745
| Std Dev. |  | 0.018348 | 0.015082 | 0.02009 | 0.012758 | 0.016903 | 0.05648 | 0.019235 | 0.021679 | 0.029665 | 0.025033 | 0.010954 | 0.018708 | 0.019408 | 0.0054772 | 0.030286 | 0.019748 | 0.029944 | 0.053048 | 0.035024 | 0.088863 | 0.015055 | 0.26009 | 0.043544 | 0.025364 | 0.008165 | 0.012111 | 0.022509 | 0.017607 | 0.02137
| Rel.Std.Dev. |  | 1.049% | 0.9183% | 1.172% | 0.7137% | 0.9901% | 3.182% | 1.036% | 1.175% | 1.735% | 1.42% | 0.6721% | 1.13% | 1.17% | 0.3391% | 1.795% | 1.145% | 1.897% | 2.973% | 1.975% | 5.56% | 0.8874% | 11.97% | 2.88% | 1.66% | 0.4929% | 0.7194% | 1.343% | 1.039% | 1.22%
| PDM3 |  | 3.086% | -3.164% | 1.072% | 5.401% | 0.6603% | 4.659% | 9.461% | 8.786% | 0.8261% | 3.971% | -3.891% | -2.417% | -2.22% | -4.775% | -0.4934% | 1.711% | -6.937% | 5.212% | 4.56% | -5.758% | 0.03993% | 28.14% | -10.84% | -9.93% | -2.319% | -0.7462% | -1.139% | -0.05834% | 3.283%
| Z-Score (Absolute) |  | 0.63 | 0.67 | 0.21 | 1.11 | 0.13 | 0.96 | 1.96 | 1.82 | 0.16 | 0.82 | 0.82 | 0.51 | 0.47 | 1 | 0.11 | 0.35 | 1.45 | 1.08 | 0.94 | 1.21 | 0 | 5.84 | 2.26 | 2.07 | 0.49 | 0.16 | 0.25 | 0.02 | 0.67
| Analytical results for Mn in OREAS 240b (Certified Value 0.079 wt.%) | [32 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.084 | 0.08065 | 0.071719 | 0.083489 | 0.080281 | 0.0849 | 0.086 | 0.0834 | 0.0816 | 0.0733 | 0.0917 | 0.0748 | 0.0738 | 0.0756 | 0.079769 | 0.0767 | 0.073 | 0.087258 | 0.0854 | 0.0753 | 0.0803 | 0.073846 | 0.078548 | 0.069 | 0.0744 | 0.0742 | 0.079 | 0.0748 | 0.0802 |  |  | 1
| 1 | 2 | 0.0851 | 0.08102 | 0.07059 | 0.082835 | 0.078888 | 0.0875 | 0.0883 | 0.0814 | 0.0829 | 0.074 | 0.092 | 0.0732 | 0.0736 | 0.0754 | 0.076849 | 0.0795 | 0.0708 | 0.083234 | 0.0817 | 0.0764 | 0.0819 | 0.074441 | 0.086067 | 0.0675 | 0.0742 | 0.0753 | 0.0783 | 0.0736 | 0.0814 |  |  | 23
| 1 | 3 | 0.0845 | 0.08236 | 0.070987 | 0.083429 | 0.079819 | 0.0831 | 0.0883 | 0.0826 | 0.0785 | 0.0733 | 0.0904 | 0.0756 | 0.0734 | 0.0761 | 0.075478 | 0.0808 | 0.0744 | 0.085297 | 0.0825 | 0.0733 | 0.0802 | 0.074599 | 0.084131 | 0.0695 | 0.0745 | 0.0744 | 0.0832 | 0.0756 | 0.0835 |  |  | 16
| 1 | 4 | 0.0849 | 0.08219 | 0.070546 | 0.08266 | 0.079769 | 0.082 | 0.0875 | 0.0848 | 0.0817 | 0.0743 | 0.0912 | 0.0735 | 0.0764 | 0.075 | 0.076124 | 0.0784 | 0.0745 | 0.084508 | 0.0846 | 0.066 | 0.0803 | 0.07457 | 0.078208 | 0.069 | 0.0749 | 0.0743 | 0.0826 | 0.0735 | 0.0818 |  |  | 0.078831
| 1 | 5 | 0.0823 | 0.08115 | 0.071014 | 0.08279 | 0.078736 | 0.0859 | 0.0875 | 0.0827 | 0.079 | 0.0738 | 0.0961 | 0.0741 | 0.0753 | 0.0748 | 0.07567 | 0.0798 | 0.0738 | 0.081682 | 0.0833 | 0.0765 | 0.0808 | 0.074414 | 0.0818 | 0.067 | 0.0747 | 0.0749 | 0.0824 | 0.0747 | 0.0807 |  |  | 104
| 1 | 6 | 0.0836 | 0.08015 | 0.069928 | 0.083169 | 0.080222 | 0.0861 | 0.0883 | 0.0827 | 0.0799 | 0.0731 | 0.0932 | 0.0739 | 0.0765 | 0.0754 | 0.076131 | 0.0792 | 0.0734 | 0.079915 | 0.0804 | 0.0745 | 0.0802 | 0.075436 | 0.079585 | 0.0655 | 0.0753 | 0.0729 | 0.0831 | 0.0747 | 0.0803
| Mean |  | 0.084067 | 0.081253 | 0.070797 | 0.083062 | 0.079619 | 0.084917 | 0.08765 | 0.082933 | 0.0806 | 0.073633 | 0.092433 | 0.074183 | 0.074833 | 0.075383 | 0.07667 | 0.079067 | 0.073317 | 0.083649 | 0.082983 | 0.073667 | 0.080617 | 0.074551 | 0.08139 | 0.067917 | 0.074667 | 0.074333 | 0.081433 | 0.074483 | 0.081317
| Median |  | 0.08425 | 0.081085 | 0.070789 | 0.083002 | 0.079794 | 0.0854 | 0.0879 | 0.0827 | 0.08075 | 0.07355 | 0.09185 | 0.074 | 0.07455 | 0.0754 | 0.076128 | 0.07935 | 0.0736 | 0.083871 | 0.0829 | 0.0749 | 0.0803 | 0.074506 | 0.080692 | 0.06825 | 0.0746 | 0.07435 | 0.0825 | 0.0747 | 0.08105
| Std Dev. |  | 0.0010289 | 0.00086586 | 0.00059905 | 0.00035088 | 0.00066005 | 0.002042 | 0.00089833 | 0.0011201 | 0.0017297 | 0.00047188 | 0.0020206 | 0.00088412 | 0.0014208 | 0.00045789 | 0.0015902 | 0.0013995 | 0.00136 | 0.0026254 | 0.0018498 | 0.0039439 | 0.00066758 | 0.0005126 | 0.0031941 | 0.0015303 | 0.00039328 | 0.0008165 | 0.0021878 | 0.00079854 | 0.0012384
| Rel.Std.Dev. |  | 1.224% | 1.066% | 0.8461% | 0.4224% | 0.829% | 2.405% | 1.025% | 1.351% | 2.146% | 0.6408% | 2.186% | 1.192% | 1.899% | 0.6074% | 2.074% | 1.77% | 1.855% | 3.139% | 2.229% | 5.354% | 0.8281% | 0.6876% | 3.924% | 2.253% | 0.5267% | 1.098% | 2.687% | 1.072% | 1.523%
| PDM3 |  | 6.642% | 3.073% | -10.19% | 5.367% | 0.9997% | 7.72% | 11.19% | 5.204% | 2.244% | -6.594% | 17.25% | -5.896% | -5.071% | -4.374% | -2.741% | 0.2989% | -6.995% | 6.112% | 5.267% | -6.551% | 2.265% | -5.429% | 3.246% | -13.85% | -5.283% | -5.706% | 3.301% | -5.515% | 3.153%
| Z-Score (Absolute) |  | 0.67 | 0.25 | 1.34 | 0.52 | 0 | 0.8 | 1.22 | 0.5 | 0.15 | 0.91 | 1.94 | 0.82 | 0.73 | 0.64 | 0.45 | 0.08 | 0.96 | 0.61 | 0.51 | 0.9 | 0.15 | 0.77 | 0.27 | 1.77 | 0.75 | 0.8 | 0.28 | 0.78 | 0.26
| Analytical results for Mo in OREAS 240b (Certified Value 2.79 ppm) | [32 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 3
| 1 | 1 | 2.86 | 2.82 | 2.77 | 2.0787 | 2.8806 | 2.77 | 2.51 | 2.7 | 3.1 | 2 | 2.9 | 2.84 | 2.69 | 2.73 | < 10 | 2.41 | 2 | < 5 | 2.1 | 2.56 | 2.88 | 1.71 | 3 | 2.8468 | 2.75 | 2.797 | 2.92 | 2.86 | 2.76 | 2.72 |  | 1
| 1 | 2 | 2.84 | 2.78 | 2.69 | 2.1095 | 2.8623 | 2.81 | 2.52 | 2.7 | 2.9 | 2 | 2.8 | 2.73 | 2.78 | 3.07 | < 10 | 2.54 | 2 | < 5 | 2.2 | 2.52 | 2.91 | 1.65 | 3 | 2.7699 | 2.74 | 2.693 | 2.95 | 2.82 | 2.78 | 2.73 |  | 24
| 1 | 3 | 2.96 | 2.81 | 2.97 | 2.102 | 2.9055 | 2.74 | 2.53 | 2.8 | 2.9 | 2 | 2.8 | 2.74 | 2.7 | 2.85 | < 10 | 2.48 | 2 | < 5 | 2.5 | 2.44 | 2.94 | 3.35 | 3 | 2.9952 | 2.81 | 2.806 | 2.88 | 3 | 2.83 | 2.79 |  | 16
| 1 | 4 | 2.83 | 2.8 | 2.84 | 2.1481 | 2.8446 | 2.63 | 2.52 | 2.8 | 3 | 2 | 2.8 | 2.88 | 2.68 | 2.7 | < 10 | 2.62 | 2 | < 5 | 2.5 | 2.2 | 3.1 | 2.61 | 3 | 2.7782 | 2.74 | 2.832 | 2.9 | 2.97 | 2.79 | 2.76 |  | 2.7855
| 1 | 5 | 2.78 | 2.82 | 4.07 | 2.0596 | 2.8663 | 2.68 | 2.42 | 2.7 | 2.7 | 2 | 3.1 | 2.74 | 2.8 | 2.76 | < 10 | 2.61 | 2 | < 5 | 2.4 | 2.55 | 2.87 | 2.64 | 3 | 2.821 | 2.76 | 2.767 | 2.83 | 2.93 | 2.79 | 2.68 |  | 105
| 1 | 6 | 2.85 | 2.74 | 2.68 | 2.0919 | 2.8064 | 2.78 | 2.47 | 2.6 | 2.7 | 2 | 2.9 | 2.88 | 2.73 | 2.82 | < 10 | 2.55 | 2 | < 5 | 2.4 | 2.44 | 2.98 | 2.3 | 3 | 2.8095 | 2.77 | 2.821 | 2.96 | 2.9 | 2.72 | 2.76
| Mean |  | 2.8533 | 2.795 | 3.0033 | 2.0983 | 2.861 | 2.735 | 2.495 | 2.7167 | 2.8833 | 2 | 2.8833 | 2.8017 | 2.73 | 2.8217 |  | 2.535 | 2 |  | 2.35 | 2.4517 | 2.9467 | 2.3767 | 3 | 2.8368 | 2.7617 | 2.786 | 2.9067 | 2.9133 | 2.7783 | 2.74
| Median |  | 2.845 | 2.805 | 2.805 | 2.0969 | 2.8643 | 2.755 | 2.515 | 2.7 | 2.9 | 2 | 2.85 | 2.79 | 2.715 | 2.79 |  | 2.545 | 2 |  | 2.4 | 2.48 | 2.925 | 2.455 | 3 | 2.8152 | 2.755 | 2.8015 | 2.91 | 2.915 | 2.785 | 2.745
| Std Dev. |  | 0.059217 | 0.030822 | 0.53343 | 0.030125 | 0.0336 | 0.06775 | 0.042308 | 0.075277 | 0.16021 | 0 | 0.1169 | 0.072778 | 0.0498 | 0.13378 |  | 0.079687 | 0 |  | 0.16432 | 0.13393 | 0.085245 | 0.64024 | 0 | 0.082583 | 0.026394 | 0.050739 | 0.048028 | 0.067429 | 0.03656 | 0.038471
| Rel.Std.Dev. |  | 2.075% | 1.103% | 17.76% | 1.436% | 1.174% | 2.477% | 1.696% | 2.771% | 5.556% | 0% | 4.054% | 2.598% | 1.824% | 4.741% |  | 3.143% | 0% |  | 6.992% | 5.463% | 2.893% | 26.94% | 0% | 2.911% | 0.9557% | 1.821% | 1.652% | 2.314% | 1.316% | 1.404%
| PDM3 |  | 2.434% | 0.3401% | 7.819% | -24.67% | 2.708% | -1.814% | -10.43% | -2.472% | 3.511% | -28.2% | 3.511% | 0.5794% | -1.993% | 1.297% |  | -8.994% | -28.2% |  | -15.64% | -11.99% | 5.785% | -14.68% | 7.7% | 1.84% | -0.8566% | 0.017% | 4.349% | 4.588% | -0.2582% | -1.634%
| Z-Score (Absolute) |  | 0.47 | 0.06 | 1.51 | 4.76 | 0.52 | 0.35 | 2.02 | 0.48 | 0.67 | NA | 0.67 | 0.11 | 0.39 | 0.25 | 15.34 | 1.74 | NA | 1.98 | 3.02 | 2.32 | 1.11 | 2.84 | NA | 0.35 | 0.17 | 0 | 0.84 | 0.88 | 0.05 | 0.32
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 240b (Certified Value 0.328 wt.%) | [32 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.34 | 0.323 | 1.2582 | 0.29952 | 0.32052 | 0.309 | 0.341 | 0.3 | 0.31 | 0.3 | 0.32 | 0.33 | 0.33 | 0.33 | 0.43881 | 0.32 | 0.31 | 0.3385 | 0.34 | 0.338 | 0.32 | 0.38 | 0.33476 | 0.3115 | 0.31 | 0.31 | 0.36 | 0.35 | 0.38 |  |  | 1
| 1 | 2 | 0.35 | 0.326 | 1.3118 | 0.29156 | 0.31392 | 0.324 | 0.341 | 0.3 | 0.32 | 0.31 | 0.33 | 0.32 | 0.32 | 0.33 | 0.42129 | 0.33 | 0.3 | 0.32395 | 0.32 | 0.341 | 0.32 | 0.38 | 0.36908 | 0.3085 | 0.31 | 0.32 | 0.38 | 0.34 | 0.39 |  |  | #N/A
| 1 | 3 | 0.35 | 0.33 | 1.162 | 0.29259 | 0.31889 | 0.302 | 0.341 | 0.3 | 0.3 | 0.3 | 0.33 | 0.33 | 0.32 | 0.33 | 0.41291 | 0.32 | 0.31 | 0.32959 | 0.32 | 0.335 | 0.31 | 0.39 | 0.35666 | 0.3165 | 0.31 | 0.31 | 0.37 | 0.35 | 0.41 |  |  | 16
| 1 | 4 | 0.35 | 0.334 | 1.2922 | 0.29466 | 0.31658 | 0.289 | 0.341 | 0.31 | 0.31 | 0.31 | 0.32 | 0.32 | 0.33 | 0.33 | 0.41895 | 0.33 | 0.31 | 0.33055 | 0.34 | 0.3 | 0.32 | 0.38 | 0.33391 | 0.314 | 0.31 | 0.31 | 0.37 | 0.34 | 0.39 |  |  | 0.3283
| 1 | 5 | 0.34 | 0.328 | 1.388 | 0.29849 | 0.31509 | 0.31 | 0.341 | 0.29 | 0.31 | 0.31 | 0.32 | 0.33 | 0.33 | 0.33 | 0.41212 | 0.33 | 0.31 | 0.31924 | 0.33 | 0.35 | 0.32 | 0.38 | 0.35351 | 0.3025 | 0.31 | 0.31 | 0.36 | 0.35 | 0.38 |  |  | 106
| 1 | 6 | 0.35 | 0.329 | 1.2575 | 0.29932 | 0.31792 | 0.31 | 0.341 | 0.29 | 0.3 | 0.3 | 0.33 | 0.33 | 0.33 | 0.33 | 0.41681 | 0.33 | 0.31 | 0.3152 | 0.34 | 0.328 | 0.32 | 0.38 | 0.3423 | 0.3 | 0.31 | 0.31 | 0.37 | 0.34 | 0.38
| Mean |  | 0.34667 | 0.32833 | 1.2783 | 0.29602 | 0.31715 | 0.30733 | 0.341 | 0.29833 | 0.30833 | 0.305 | 0.325 | 0.32667 | 0.32667 | 0.33 | 0.42015 | 0.32667 | 0.30833 | 0.32617 | 0.33167 | 0.332 | 0.31833 | 0.38167 | 0.34837 | 0.30883 | 0.31 | 0.31167 | 0.36833 | 0.345 | 0.38833
| Median |  | 0.35 | 0.3285 | 1.2752 | 0.29657 | 0.31725 | 0.3095 | 0.341 | 0.3 | 0.31 | 0.305 | 0.325 | 0.33 | 0.33 | 0.33 | 0.41788 | 0.33 | 0.31 | 0.32677 | 0.335 | 0.3365 | 0.32 | 0.38 | 0.3479 | 0.31 | 0.31 | 0.31 | 0.37 | 0.345 | 0.385
| Std Dev. |  | 0.005164 | 0.0037238 | 0.074448 | 0.0035425 | 0.0024509 | 0.011483 | 0 | 0.0075277 | 0.0075277 | 0.0054772 | 0.0054772 | 0.005164 | 0.005164 | 0 | 0.0097874 | 0.005164 | 0.0040825 | 0.0084428 | 0.0098319 | 0.017263 | 0.0040825 | 0.0040825 | 0.013824 | 0.0064936 | 0 | 0.0040825 | 0.0075277 | 0.0054772 | 0.01169
| Rel.Std.Dev. |  | 1.49% | 1.134% | 5.824% | 1.197% | 0.7728% | 3.736% | 0% | 2.523% | 2.441% | 1.796% | 1.685% | 1.581% | 1.581% | 0% | 2.33% | 1.581% | 1.324% | 2.588% | 2.964% | 5.2% | 1.282% | 1.07% | 3.968% | 2.103% | 0% | 1.31% | 2.044% | 1.588% | 3.01%
| PDM3 |  | 5.593% | 0.009069% | 289.4% | -9.832% | -3.397% | -6.387% | 3.867% | -9.129% | -6.083% | -7.098% | -1.006% | -0.4986% | -0.4986% | 0.5167% | 27.98% | -0.4986% | -6.083% | -0.6495% | 1.024% | 1.126% | -3.037% | 16.25% | 6.112% | -5.931% | -5.575% | -5.068% | 12.19% | 5.086% | 18.28%
| Z-Score (Absolute) |  | 0.76 | 0.06 | 35.98 | 1.16 | 0.36 | 0.73 | 0.55 | 1.07 | 0.69 | 0.82 | 0.06 | 0 | 0 | 0.13 | 3.53 | 0 | 0.69 | 0.02 | 0.19 | 0.2 | 0.32 | 2.08 | 0.82 | 0.67 | 0.63 | 0.57 | 1.58 | 0.69 | 2.33
| Analytical results for Nb in OREAS 240b (Indicative Value 0.11 ppm) | [32 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | < 0.05 | 0.13 | 33.89 | 0.13569 | < 0.1 | 0.34 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.1 | 0.12 | 0.09 | 0.06 | < 10 | < 0.2 | < 1 | 0.12 | 0.38 | 0.2 | 0.07 | 0.13 | [9 empty] | 1
| 1 | 2 | < 0.05 | 0.12 | 34.3 | 0.12822 | < 0.1 | 0.36 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.09 | 0.11 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.28 | 0.19 | 0.07 | 0.13 | [9 empty] | 7
| 1 | 3 | < 0.05 | 0.12 | 33.59 | 0.13725 | < 0.1 | 0.31 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.09 | 0.12 | 0.08 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.21 | 0.19 | 0.08 | 0.12 | [9 empty] | 16
| 1 | 4 | < 0.05 | 0.12 | 32.7 | 0.1656 | < 0.1 | 0.3 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.11 | 0.12 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | 0.15 | 0.21 | 0.08 | 0.13 | [9 empty] | 0.11255
| 1 | 5 | < 0.05 | 0.12 | 40.8 | 0.14922 | < 0.1 | 0.33 | < 0.2 | 0.07 | < 5 | < 0.1 | 0.09 | 0.13 | 0.08 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | < 0.05 | 0.2 | 0.08 | 0.14 | [9 empty] | 13
| 1 | 6 | < 0.05 | 0.12 | 32.89 | 0.15263 | < 0.1 | 0.3 | < 0.2 | 0.07 | < 5 | < 0.1 | 0.09 | 0.13 | 0.09 | 0.07 | < 10 | < 0.2 | < 1 | 0.11 | < 0.05 | 0.19 | 0.07 | 0.14
| Mean |  |  | 0.12167 | 34.695 | 0.14477 |  | 0.32333 |  | 0.085 |  |  | 0.095 | 0.12167 | 0.086667 | 0.068333 |  |  |  | 0.11167 | 0.255 | 0.19667 | 0.075 | 0.13167
| Median |  |  | 0.12 | 33.74 | 0.14323 |  | 0.32 |  | 0.09 |  |  | 0.09 | 0.12 | 0.09 | 0.07 |  |  |  | 0.11 | 0.245 | 0.195 | 0.075 | 0.13
| Std Dev. |  |  | 0.0040825 | 3.0507 | 0.013624 |  | 0.024221 |  | 0.012247 |  |  | 0.0083666 | 0.0075277 | 0.005164 | 0.0040825 |  |  |  | 0.0040825 | 0.098826 | 0.008165 | 0.0054772 | 0.0075277
| Rel.Std.Dev. |  |  | 3.355% | 8.793% | 9.411% |  | 7.491% |  | 14.41% |  |  | 8.807% | 6.187% | 5.958% | 5.974% |  |  |  | 3.656% | 38.76% | 4.152% | 7.303% | 5.717%
| PDM3 |  |  | 8.096% | 30725% | 28.62% |  | 187.3% |  | -24.48% |  |  | -15.6% | 8.096% | -23% | -39.29% |  |  |  | -0.7889% | 126.6% | 74.73% | -33.37% | 16.98%
| Z-Score (Absolute) |  | 1.37 | 0.07 | 518.14 | 0.42 | 1 | 3.1 | 0.25 | 0.47 | 35.71 | 1 | 0.32 | 0.07 | 0.45 | 0.72 | 73.17 | 0.25 | 5.74 | 0.07 | 0.92 | 1.2 | 0.62 | 0.22
| Analytical results for Nd in OREAS 240b (Certified Value 6.26 ppm) | [32 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 6.97 | 6 | 5.2182 | 6.32 | 6.66 | 6 | [25 empty] | 1
| 1 | 2 | 6.906 | 5.953 | 5.2388 | 6.47 | 6.79 | 6 | [25 empty] | 26
| 1 | 3 | 6.945 | 6.176 | 5.2816 | 6.65 | 6.66 | 6 | [25 empty] | 16
| 1 | 4 | 7.042 | 5.976 | 5.2778 | 6.54 | 6.69 | 6.5 | [25 empty] | 6.2576
| 1 | 5 | 7.023 | 6.233 | 5.2777 | 6.53 | 6.64 | 6 | [25 empty] | 107
| 1 | 6 | 6.823 | 6.524 | 5.2574 | 6.74 | 6.46 | 6
| Mean |  | 6.9515 | 6.1437 | 5.2586 | 6.5417 | 6.65 | 6.0833
| Median |  | 6.9575 | 6.088 | 5.2676 | 6.535 | 6.66 | 6
| Std Dev. |  | 0.080336 | 0.21853 | 0.025645 | 0.14525 | 0.10733 | 0.20412
| Rel.Std.Dev. |  | 1.156% | 3.557% | 0.4877% | 2.22% | 1.614% | 3.355%
| PDM3 |  | 11.09% | -1.82% | -15.96% | 4.54% | 6.271% | -2.784%
| Z-Score (Absolute) |  | 1.45 | 0.47 | 2.58 | 0.47 | 0.73 | 0.62
| Analytical results for Ni in OREAS 240b (Certified Value 68 ppm) | [32 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 1
| 1 | 1 | 68.2 | 66.8 | 60.37 | 66.13 | 71.8212 | 67.2 | 69.2 | 74 | 70.3 | 61 | 61.9 | 69.4 | 67.5 | 70.5 | 72.155 | 64.5 | 59 | 73.9922 | 71 | 71.1 | 68.2 | 65.79 | 63.2408 | 63 | 72.1033 | 67.4 | 66.1 | 69.8 | 63.2 | 67.5 |  | 1
| 1 | 2 | 69.6 | 68.3 | 59.17 | 65.31 | 70.2878 | 69.5 | 70 | 73 | 71.4 | 62 | 61.7 | 68.7 | 67.6 | 71.5 | 68.037 | 67.5 | 57 | 72.0636 | 71 | 68.8 | 71.9 | 64.53 | 65.1176 | 62 | 73.231 | 66.5 | 67.5 | 70.9 | 63.2 | 67 |  | 27
| 1 | 3 | 69.3 | 68.4 | 59.65 | 65.04 | 69.8237 | 67.6 | 68.6 | 73 | 68.2 | 61 | 61.4 | 67.9 | 67.7 | 70 | 66.818 | 66.8 | 60 | 74.1946 | 69 | 69.4 | 71 | 65.36 | 65.458 | 65 | 73.1875 | 66.7 | 68.4 | 73.9 | 64.6 | 68.9 |  | 16
| 1 | 4 | 69.3 | 68.4 | 59.74 | 65.88 | 70.0186 | 66.3 | 70.8 | 73 | 70.9 | 61 | 62.2 | 69.6 | 67.9 | 68.6 | 69.178 | 69 | 60 | 72.7351 | 69 | 62.2 | 71.8 | 65.35 | 63.1672 | 64 | 70.1516 | 67.5 | 66.8 | 71.8 | 62.4 | 69.4 |  | 67.9061
| 1 | 5 | 67.3 | 69.1 | 58.5 | 66.13 | 69.359 | 69.2 | 67.8 | 71 | 68 | 62 | 64.8 | 69.1 | 67.8 | 69.6 | 67.657 | 68.7 | 59 | 70.8509 | 67 | 71.4 | 70.2 | 64.77 | 67.114 | 61 | 73.6061 | 66.8 | 66.8 | 72.1 | 64.2 | 67.5 |  | 108
| 1 | 6 | 68.4 | 66.5 | 58.55 | 66.95 | 70.548 | 69.5 | 69.5 | 72 | 66.9 | 61 | 61.9 | 69.5 | 67.5 | 70.1 | 67.699 | 68.8 | 59 | 69.513 | 67 | 69.6 | 68.1 | 66.69 | 63.9492 | 60 | 70.8064 | 67.3 | 65.8 | 73.1 | 62.6 | 69.2
| Mean |  | 68.6833 | 67.9167 | 59.33 | 65.9067 | 70.3097 | 68.2167 | 69.3167 | 72.6667 | 69.2833 | 61.3333 | 62.3167 | 69.0333 | 67.6667 | 70.05 | 68.5907 | 67.55 | 59 | 72.2249 | 69 | 68.75 | 70.2 | 65.415 | 64.6745 | 62.5 | 72.181 | 67.0333 | 66.9 | 71.9333 | 63.3667 | 68.25
| Median |  | 68.85 | 68.35 | 59.41 | 66.005 | 70.1532 | 68.4 | 69.35 | 73 | 69.25 | 61 | 61.9 | 69.25 | 67.65 | 70.05 | 67.868 | 68.1 | 59 | 72.3994 | 69 | 69.5 | 70.6 | 65.355 | 64.5334 | 62.5 | 72.6454 | 67.05 | 66.8 | 71.95 | 63.2 | 68.2
| Std Dev. |  | 0.87502 | 1.0265 | 0.73144 | 0.67766 | 0.84462 | 1.3674 | 1.0515 | 1.0328 | 1.8236 | 0.5164 | 1.2449 | 0.64395 | 0.1633 | 0.96073 | 1.906 | 1.7237 | 1.0954 | 1.8183 | 1.7889 | 3.3644 | 1.7029 | 0.77107 | 1.5241 | 1.8708 | 1.4255 | 0.41793 | 0.94657 | 1.476 | 0.87101 | 1.033
| Rel.Std.Dev. |  | 1.274% | 1.511% | 1.233% | 1.028% | 1.201% | 2.004% | 1.517% | 1.421% | 2.632% | 0.842% | 1.998% | 0.9328% | 0.2413% | 1.371% | 2.779% | 2.552% | 1.857% | 2.518% | 2.593% | 4.894% | 2.426% | 1.179% | 2.357% | 2.993% | 1.975% | 0.6235% | 1.415% | 2.052% | 1.375% | 1.513%
| PDM3 |  | 1.145% | 0.01563% | -12.63% | -2.944% | 3.54% | 0.4574% | 2.077% | 7.011% | 2.028% | -9.679% | -8.231% | 1.66% | -0.3525% | 3.157% | 1.008% | -0.5243% | -13.12% | 6.36% | 1.611% | 1.243% | 3.378% | -3.668% | -4.759% | -7.961% | 6.295% | -1.285% | -1.482% | 5.931% | -6.685% | 0.5065%
| Z-Score (Absolute) |  | 0.16 | 0.11 | 3.17 | 0.83 | 0.74 | 0.01 | 0.39 | 1.58 | 0.37 | 2.46 | 2.11 | 0.29 | 0.2 | 0.65 | 0.13 | 0.24 | 3.29 | 1.42 | 0.27 | 0.18 | 0.7 | 1 | 1.27 | 2.04 | 1.41 | 0.43 | 0.48 | 1.32 | 1.73 | 0.01
| Analytical results for P in OREAS 240b (Certified Value 0.042 wt.%) | [32 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.0433 | 0.0422 | 0.035691 | 0.042322 | 0.039 | 0.0406 | 0.0427 | 0.0408 | 0.0422 | 0.045 | 0.042 | 0.042 | 0.04 | 0.047443 | 0.041 | 0.0383 | 0.045691 | 0.043 | 0.041 | 0.04 | 0.0405 | 0.041 | 0.041 | 0.0431 | 0.0409 | 0.042 | [5 empty] | 1
| 1 | 2 | 0.0443 | 0.043 | 0.035618 | 0.041692 | 0.039 | 0.0401 | 0.0418 | 0.0397 | 0.0425 | 0.046 | 0.041 | 0.041 | 0.039 | 0.045489 | 0.042 | 0.0368 | 0.044347 | 0.041 | 0.041 | 0.04 | 0.041 | 0.041 | 0.041 | 0.0433 | 0.0402 | 0.042 | [5 empty] | #N/A
| 1 | 3 | 0.044 | 0.0426 | 0.035596 | 0.041697 | 0.039 | 0.041 | 0.0424 | 0.0373 | 0.0423 | 0.044 | 0.042 | 0.042 | 0.04 | 0.044277 | 0.042 | 0.0385 | 0.045597 | 0.043 | 0.04 | 0.04 | 0.041 | 0.041 | 0.042 | 0.0439 | 0.0415 | 0.043 | [5 empty] | 16
| 1 | 4 | 0.0439 | 0.0421 | 0.035316 | 0.041122 | 0.038 | 0.0406 | 0.043 | 0.0398 | 0.0427 | 0.044 | 0.041 | 0.043 | 0.039 | 0.045064 | 0.042 | 0.0382 | 0.043811 | 0.044 | 0.036 | 0.04 | 0.041 | 0.041 | 0.041 | 0.0441 | 0.0403 | 0.042 | [5 empty] | 0.041627
| 1 | 5 | 0.0429 | 0.0425 | 0.035577 | 0.041274 | 0.04 | 0.0406 | 0.0421 | 0.0387 | 0.0426 | 0.047 | 0.042 | 0.042 | 0.039 | 0.045216 | 0.042 | 0.0384 | 0.042888 | 0.043 | 0.041 | 0.04 | 0.0395 | 0.041 | 0.043 | 0.0437 | 0.0407 | 0.041 | [5 empty] | 109
| 1 | 6 | 0.0434 | 0.0414 | 0.035402 | 0.041812 | 0.039 | 0.0393 | 0.0422 | 0.0403 | 0.0419 | 0.044 | 0.042 | 0.043 | 0.039 | 0.044564 | 0.042 | 0.0384 | 0.042534 | 0.041 | 0.039 | 0.04 | 0.0395 | 0.041 | 0.041 | 0.0435 | 0.0407 | 0.042
| Mean |  | 0.043633 | 0.0423 | 0.035533 | 0.041653 | 0.039 | 0.040367 | 0.042367 | 0.039433 | 0.042367 | 0.045 | 0.041667 | 0.042167 | 0.039333 | 0.045342 | 0.041833 | 0.0381 | 0.044145 | 0.0425 | 0.039667 | 0.04 | 0.040417 | 0.041 | 0.0415 | 0.0436 | 0.040717 | 0.042
| Median |  | 0.04365 | 0.04235 | 0.035587 | 0.041694 | 0.039 | 0.0406 | 0.0423 | 0.03975 | 0.0424 | 0.0445 | 0.042 | 0.042 | 0.039 | 0.04514 | 0.042 | 0.03835 | 0.044079 | 0.043 | 0.0405 | 0.04 | 0.04075 | 0.041 | 0.041 | 0.0436 | 0.0407 | 0.042
| Std Dev. |  | 0.00052026 | 0.00054406 | 0.00014307 | 0.00042467 | 0.00063246 | 0.00059554 | 0.00043205 | 0.001258 | 0.00029439 | 0.0012649 | 0.0005164 | 0.00075277 | 0.0005164 | 0.0011196 | 0.00040825 | 0.00064498 | 0.0013284 | 0.0012247 | 0.0019664 | 0 | 0.00073598 | 0 | 0.00083666 | 0.00037417 | 0.00046655 | 0.00063246
| Rel.Std.Dev. |  | 1.192% | 1.286% | 0.4026% | 1.02% | 1.622% | 1.475% | 1.02% | 3.19% | 0.6949% | 2.811% | 1.239% | 1.785% | 1.313% | 2.469% | 0.9759% | 1.693% | 3.009% | 2.882% | 4.957% | 0% | 1.821% | 0% | 2.016% | 0.8582% | 1.146% | 1.506%
| PDM3 |  | 4.82% | 1.616% | -14.64% | 0.06202% | -6.311% | -3.028% | 1.777% | -5.27% | 1.777% | 8.103% | 0.09503% | 1.296% | -5.51% | 8.925% | 0.4954% | -8.473% | 6.048% | 2.097% | -4.71% | -3.909% | -2.908% | -1.506% | -0.3054% | 4.739% | -2.187% | 0.8958%
| Z-Score (Absolute) |  | 1.05 | 0.34 | 3.26 | 0 | 1.41 | 0.69 | 0.38 | 1.18 | 0.38 | 1.77 | 0 | 0.27 | 1.24 | 1.96 | 0.09 | 1.89 | 1.32 | 0.45 | 1.06 | 0.88 | 0.66 | 0.35 | 0.08 | 1.03 | 0.5 | 0.18
| Analytical results for Pb in OREAS 240b (Certified Value 40.7 ppm) | [32 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 2
| 1 | 1 | 43.9 | 40.9 | 35.13 | 39.85 | 39.7974 | 31 | 41.1 | 37.8 | 41.6 | 38 | 44.5 | 40.6 | 39.4 | 43 | 39.1 | 35 | 39.5094 | 34.6 | 39.41 | 41.4 | 46.35 | 39.2088 | 43 | 41.1299 | 41.2 | 48.738 | 40.5 | 38.9 | 39.9 | 41.5 |  | 1
| 1 | 2 | 45.2 | 40.2 | 38.88 | 38.6 | 39.5029 | 31.4 | 41.8 | 36.8 | 42.1 | 38 | 44.7 | 40.6 | 40.1 | 44.5 | 41 | 33 | 38.1526 | 34 | 38.79 | 41.8 | 42.23 | 43.0311 | 44 | 42.2561 | 40.7 | 50.601 | 40.7 | 39 | 40.7 | 42.2 |  | 28
| 1 | 3 | 44.5 | 40.1 | 34.09 | 39.49 | 38.9658 | 30.5 | 41.9 | 38.4 | 40.9 | 38 | 44.7 | 39.2 | 39.1 | 43.4 | 40 | 35 | 40.0098 | 33.3 | 36.92 | 41.5 | 41.38 | 41.199 | 43 | 42.5173 | 40.3 | 55.738 | 40.8 | 40.5 | 41.7 | 42.3 |  | 16
| 1 | 4 | 44.4 | 41 | 35.6 | 38.91 | 38.9431 | 30.8 | 41.6 | 38.7 | 41.5 | 38 | 45.1 | 40.7 | 39.4 | 42.9 | 41.4 | 35 | 39.9235 | 33.1 | 33.48 | 42 | 41.34 | 40.083 | 44 | 42.9287 | 40.3 | 55.54 | 41 | 39.8 | 40.2 | 42.6 |  | 40.6848
| 1 | 5 | 43.9 | 40.7 | 33.83 | 38.84 | 39.0262 | 31.5 | 40.3 | 37 | 41.4 | 39 | 47.9 | 40.1 | 39.3 | 42.6 | 41.2 | 35 | 37.8932 | 32.5 | 38.44 | 40.8 | 43.98 | 40.548 | 42 | 41.2678 | 40.4 | 47.277 | 40.2 | 41.4 | 40.8 | 41.7 |  | 110
| 1 | 6 | 44.3 | 40.2 | 36.98 | 39.44 | 39.2575 | 31.2 | 42.3 | 37.5 | 40.5 | 38 | 45.5 | 39.9 | 39.3 | 41.7 | 41.4 | 35 | 37.4837 | 34.9 | 37.87 | 42.8 | 42.34 | 40.3713 | 42 | 42.418 | 40.8 | 50.874 | 40.6 | 40.1 | 41.7 | 42.4
| Mean |  | 44.3667 | 40.5167 | 35.7517 | 39.1883 | 39.2488 | 31.0667 | 41.5 | 37.7 | 41.3333 | 38.1667 | 45.4 | 40.1833 | 39.4333 | 43.0167 | 40.6833 | 34.6667 | 38.8287 | 33.7333 | 37.485 | 41.7167 | 42.9367 | 40.7402 | 43 | 42.0863 | 40.6167 | 51.4613 | 40.6333 | 39.95 | 40.8333 | 42.1167
| Median |  | 44.35 | 40.45 | 35.365 | 39.175 | 39.1419 | 31.1 | 41.7 | 37.65 | 41.45 | 38 | 44.9 | 40.35 | 39.35 | 42.95 | 41.1 | 35 | 38.831 | 33.65 | 38.155 | 41.65 | 42.285 | 40.4596 | 43 | 42.337 | 40.55 | 50.7375 | 40.65 | 39.95 | 40.75 | 42.25
| Std Dev. |  | 0.48028 | 0.39707 | 1.9057 | 0.47688 | 0.34313 | 0.37771 | 0.70711 | 0.75366 | 0.56095 | 0.40825 | 1.2759 | 0.57764 | 0.34448 | 0.92394 | 0.9347 | 0.8165 | 1.1134 | 0.92664 | 2.1367 | 0.67057 | 1.927 | 1.2968 | 0.89443 | 0.72373 | 0.35449 | 3.491 | 0.27325 | 0.94393 | 0.74744 | 0.42622
| Rel.Std.Dev. |  | 1.083% | 0.98% | 5.33% | 1.217% | 0.8743% | 1.216% | 1.704% | 1.999% | 1.357% | 1.07% | 2.81% | 1.438% | 0.8736% | 2.148% | 2.298% | 2.355% | 2.867% | 2.747% | 5.7% | 1.607% | 4.488% | 3.183% | 2.08% | 1.72% | 0.8728% | 6.784% | 0.6725% | 2.363% | 1.83% | 1.012%
| PDM3 |  | 9.05% | -0.4133% | -12.13% | -3.678% | -3.53% | -23.64% | 2.004% | -7.336% | 1.594% | -6.189% | 11.59% | -1.233% | -3.076% | 5.732% | -0.003624% | -14.79% | -4.562% | -17.09% | -7.865% | 2.536% | 5.535% | 0.1361% | 5.691% | 3.445% | -0.1675% | 26.49% | -0.1265% | -1.806% | 0.3651% | 3.519%
| Z-Score (Absolute) |  | 1.71 | 0.05 | 2.23 | 0.66 | 0.63 | 4.37 | 0.4 | 1.34 | 0.32 | 1.12 | 2.18 | 0.2 | 0.54 | 1.09 | 0.03 | 2.72 | 0.82 | 3.15 | 1.43 | 0.5 | 1.06 | 0.05 | 1.08 | 0.67 | 0 | 4.95 | 0 | 0.31 | 0.1 | 0.68
| Analytical results for Pd in OREAS 240b (Indicative Value 12.5 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 1
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 1
| 1 | 1 | 14 | 16 | 13 | < 50 | 26.02 | 24 | [25 empty] | 1
| 1 | 2 | 14 | 11 | 10 | < 50 | < 20 | 28 | [25 empty] | 1
| 1 | 3 | 13 | 18 | 14 | < 50 | 25.87 | 29 | [25 empty] | 16
| 1 | 4 | 12 | 11 | 14 | < 50 | < 20 | 30 | [25 empty] | 12.5
| 1 | 5 | 13 | 15 | 13 | < 50 | < 20 | 34 | [25 empty] | 11
| 1 | 6 | 13 | 16 | 10 | < 50 | < 20 | 29
| Mean |  | 13.1667 | 14.5 | 12.3333 |  | 25.945 | 29
| Median |  | 13 | 15.5 | 13 |  | 25.945 | 29
| Std Dev. |  | 0.75277 | 2.881 | 1.8619 |  | 0.10607 | 3.2249
| Rel.Std.Dev. |  | 5.717% | 19.87% | 15.1% |  | 0.4088% | 11.12%
| PDM3 |  | 5.333% | 16% | -1.333% |  | 107.6% | 132%
| Z-Score (Absolute) |  | 0.54 | 0.13 | 0.8 | 3.15 | 0.13 | 4.4
| Analytical results for Pr in OREAS 240b (Certified Value 1.39 ppm) | [32 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 1.497 | 1.317 | 1.2071 | 1.4 | 1.43 | 1.4 | [25 empty] | 1
| 1 | 2 | 1.496 | 1.384 | 1.2239 | 1.5 | 1.45 | 1.4 | [25 empty] | 29
| 1 | 3 | 1.492 | 1.35 | 1.1991 | 1.5 | 1.44 | 1.4 | [25 empty] | 16
| 1 | 4 | 1.489 | 1.332 | 1.1976 | 1.4 | 1.44 | 1.5 | [25 empty] | 1.392
| 1 | 5 | 1.516 | 1.277 | 1.2044 | 1.5 | 1.42 | 1.4 | [25 empty] | 111
| 1 | 6 | 1.476 | 1.364 | 1.1994 | 1.6 | 1.41 | 1.3
| Mean |  | 1.4943 | 1.3373 | 1.2052 | 1.4833 | 1.4317 | 1.4
| Median |  | 1.494 | 1.341 | 1.2019 | 1.5 | 1.435 | 1.4
| Std Dev. |  | 0.013033 | 0.037766 | 0.0098115 | 0.075277 | 0.01472 | 0.063246
| Rel.Std.Dev. |  | 0.8722% | 2.824% | 0.8141% | 5.075% | 1.028% | 4.518%
| PDM3 |  | 7.353% | -3.926% | -13.42% | 6.563% | 2.851% | 0.5759%
| Z-Score (Absolute) |  | 0.73 | 0.73 | 1.95 | 0.62 | 0.15 | 0.15
| Analytical results for Pt in OREAS 240b (Indicative Value 6.8 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 2
| 1 | 1 | 7 | < 5 | < 5 | < 10 | 5.44 | < 10 | [25 empty] | 1
| 1 | 2 | 7 | 6 | < 5 | < 10 | 7.82 | < 10 | [25 empty] | 1
| 1 | 3 | 6 | 7 | < 5 | < 10 | 5.36 | < 10 | [25 empty] | 16
| 1 | 4 | 7 | 9 | < 5 | < 10 | 9.88 | < 10 | [25 empty] | 6.7994
| 1 | 5 | 8 | 8 | < 5 | < 10 | 9.01 | < 10 | [25 empty] | 12
| 1 | 6 | 6 | < 5 | < 5 | < 10 | 8.88 | < 10
| Mean |  | 6.8333 | 7.5 |  |  | 7.7317
| Median |  | 7 | 7.5 |  |  | 8.35
| Std Dev. |  | 0.75277 | 1.291 |  |  | 1.9211
| Rel.Std.Dev. |  | 11.02% | 17.21% |  |  | 24.85%
| PDM3 |  | 0.4984% | 10.3% |  |  | 13.71%
| Z-Score (Absolute) |  | 1.04 | 0.31 | 2.15 | 0.31 | 1.7 | 0.31
| Analytical results for Rb in OREAS 240b (Certified Value 8.91 ppm) | [32 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 8.8 | 9.14 | 8.0388 | 8.8657 | 7.2 | 10.5 | 8.93 | 9.12 | 9.9 | 8.4 | 8.7 | 8.3 | 8.1 | 8.46 | 9.2 | 8.5 | 9 | 8.8 | 8.8 | 9.4 | 8.4 | [10 empty] | 1
| 1 | 2 | 8.9 | 8.99 | 8.0058 | 8.6894 | 7.4 | 10.1 | 8.91 | 9.23 | 10.2 | 8.3 | 8.7 | 8.2 | 8.5 | 8.91 | 9.5 | 8.5 | 9 | 8.5 | 8.6 | 9.3 | 8.2 | [10 empty] | 30
| 1 | 3 | 8.8 | 9.05 | 7.8999 | 8.8068 | 7.1 | 10.3 | 9.06 | 8.92 | 10.2 | 8.1 | 8.8 | 8.1 | 8.4 | 8.89 | 9 | 8.5 | 9.2 | 8.9 | 9.1 | 9.6 | 8.5 | [10 empty] | 16
| 1 | 4 | 8.8 | 9.2 | 7.9421 | 8.685 | 6.9 | 10.2 | 9.59 | 9.07 | 10.2 | 8.4 | 9 | 8 | 8.7 | 8.94 | 9.6 | 8.5 | 9.1 | 8.8 | 9.1 | 9.1 | 8.5 | [10 empty] | 8.9095
| 1 | 5 | 8.7 | 9.13 | 7.794 | 8.691 | 7.2 | 9.72 | 9.87 | 9.27 | 10.6 | 8.2 | 8.8 | 8.2 | 8.7 | 8.94 | 9.3 | 8 | 9.4 | 8.7 | 9 | 9.4 | 8.3 | [10 empty] | 112
| 1 | 6 | 8.8 | 9.01 | 7.8627 | 8.8783 | 7.5 | 9.91 | 9.96 | 9.07 | 10 | 8.3 | 8.8 | 8.2 | 8.6 | 8.39 | 9.6 | 8 | 9.3 | 8.8 | 8.9 | 9.1 | 8.5
| Mean |  | 8.8 | 9.0867 | 7.9239 | 8.7694 | 7.2167 | 10.1217 | 9.3867 | 9.1133 | 10.1833 | 8.2833 | 8.8 | 8.1667 | 8.5 | 8.755 | 9.3667 | 8.3333 | 9.1667 | 8.75 | 8.9167 | 9.3167 | 8.4
| Median |  | 8.8 | 9.09 | 7.921 | 8.7489 | 7.2 | 10.15 | 9.325 | 9.095 | 10.2 | 8.3 | 8.8 | 8.2 | 8.55 | 8.9 | 9.4 | 8.5 | 9.15 | 8.8 | 8.95 | 9.35 | 8.45
| Std Dev. |  | 0.063246 | 0.082624 | 0.091047 | 0.09188 | 0.2137 | 0.27845 | 0.47878 | 0.12596 | 0.24014 | 0.1169 | 0.10954 | 0.10328 | 0.22804 | 0.25727 | 0.24221 | 0.2582 | 0.1633 | 0.13784 | 0.19408 | 0.19408 | 0.12649
| Rel.Std.Dev. |  | 0.7187% | 0.9093% | 1.149% | 1.048% | 2.961% | 2.751% | 5.101% | 1.382% | 2.358% | 1.411% | 1.245% | 1.265% | 2.683% | 2.939% | 2.586% | 3.098% | 1.781% | 1.575% | 2.177% | 2.083% | 1.506%
| PDM3 |  | -1.229% | 1.989% | -11.06% | -1.572% | -19% | 13.61% | 5.356% | 2.288% | 14.3% | -7.028% | -1.229% | -8.338% | -4.596% | -1.734% | 5.131% | -6.467% | 2.886% | -1.79% | 0.08046% | 4.57% | -5.719%
| Z-Score (Absolute) |  | 0 | 0.48 | 1.48 | 0.05 | 2.67 | 2.23 | 0.99 | 0.53 | 2.33 | 0.87 | 0 | 1.07 | 0.51 | 0.08 | 0.96 | 0.79 | 0.62 | 0.08 | 0.2 | 0.87 | 0.67
| Analytical results for Re in OREAS 240b (Certified Value 0.002 ppm) | [32 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | 0.005 | 0.004 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.007 | < 0.002 | 0.003 | 0.004 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.001 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.01 | < 0.002 | 0.003 | 0.004 | [13 empty] | 31
| 1 | 3 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.001 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.012 | < 0.002 | 0.003 | 0.001 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.001 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.011 | < 0.002 | 0.003 | < 0.001 | [13 empty] | 0.0019375
| 1 | 5 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | 0.005 | 0.001 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.012 | < 0.002 | 0.003 | 0.003 | [13 empty] | 113
| 1 | 6 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.01 | < 0.002 | 0.002 | < 0.001
| Mean |  |  | 0.002 | 0.002 |  | 0.002 | 0.005 | 0.0023333 | 0.002 | 0.0018333 | 0.0018333 |  |  |  | 0.0018333 | 0.010333 |  | 0.0028333 | 0.003
| Median |  |  | 0.002 | 0.002 |  | 0.002 | 0.005 | 0.002 | 0.002 | 0.002 | 0.002 |  |  |  | 0.002 | 0.0105 |  | 0.003 | 0.0035
| Std Dev. |  |  | 0 | 0 |  | 0.00063246 | 0 | 0.0010328 | 0 | 0.00040825 | 0.00040825 |  |  |  | 0.00040825 | 0.0018619 |  | 0.00040825 | 0.0014142
| Rel.Std.Dev. |  |  | 0% | 0% |  | 31.62% | 0% | 44.26% | 0% | 22.27% | 22.27% |  |  |  | 22.27% | 18.02% |  | 14.41% | 47.14%
| PDM3 |  |  | 3.226% | 3.226% |  | 3.226% | 158.1% | 20.43% | 3.226% | -5.376% | -5.376% |  |  |  | -5.376% | 433.3% |  | 46.24% | 54.84%
| Z-Score (Absolute) |  | 64.4 | 0.34 | 0.34 | 30.68 | 0.34 | 1.46 | 0.11 | 0.34 | 0.56 | 0.56 | 3368.51 | 30.68 | 30.68 | 0.56 | 10.9 | 1.69 | 0.79 | 0.11
| Analytical results for S in OREAS 240b (Certified Value 0.545 wt.%) | [32 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.56 | 0.53 | 0.44536 | 0.562 | 0.56745 | 0.521 | 0.548 | 0.54 | 0.52 | 0.56 | 0.54 | 0.56 | 0.53 | 0.55 | 0.55 | 0.48 | 0.51001 | 0.56 | 0.49 | 0.55 | 0.52836 | 0.5415 | 0.52 | 0.55 | 0.58 | 0.54 | 0.56 |  |  |  |  | 1
| 1 | 2 | 0.57 | 0.53 | 0.49212 | 0.56761 | 0.55395 | 0.526 | 0.547 | 0.54 | 0.54 | 0.56 | 0.55 | 0.55 | 0.55 | 0.54 | 0.56 | 0.46 | 0.50057 | 0.55 | 0.49 | 0.56 | 0.52565 | 0.5385 | 0.52 | 0.53 | 0.57 | 0.53 | 0.56 |  |  |  |  | 16
| 1 | 3 | 0.56 | 0.53 | 0.43786 | 0.56274 | 0.56133 | 0.523 | 0.553 | 0.56 | 0.52 | 0.56 | 0.54 | 0.56 | 0.54 | 0.54 | 0.57 | 0.48 | 0.52209 | 0.53 | 0.48 | 0.55 | 0.5286 | 0.547 | 0.52 | 0.55 | 0.58 | 0.55 | 0.58 |  |  |  |  | 16
| 1 | 4 | 0.57 | 0.54 | 0.45566 | 0.56641 | 0.55366 | 0.522 | 0.546 | 0.56 | 0.56 | 0.56 | 0.55 | 0.55 | 0.55 | 0.54 | 0.56 | 0.48 | 0.50629 | 0.56 | 0.42 | 0.55 | 0.53005 | 0.5425 | 0.52 | 0.53 | 0.57 | 0.53 | 0.56 |  |  |  |  | 0.54527
| 1 | 5 | 0.56 | 0.53 | 0.39959 | 0.56555 | 0.56124 | 0.533 | 0.543 | 0.54 | 0.54 | 0.55 | 0.53 | 0.56 | 0.55 | 0.55 | 0.56 | 0.48 | 0.49719 | 0.53 | 0.49 | 0.56 | 0.55587 | 0.532 | 0.52 | 0.55 | 0.58 | 0.54 | 0.56 |  |  |  |  | 114
| 1 | 6 | 0.56 | 0.53 | 0.4761 | 0.56316 | 0.56068 | 0.528 | 0.546 | 0.54 | 0.54 | 0.54 | 0.54 | 0.55 | 0.54 | 0.54 | 0.56 | 0.48 | 0.48598 | 0.52 | 0.47 | 0.56 | 0.53841 | 0.5305 | 0.52 | 0.54 | 0.58 | 0.54 | 0.56
| Mean |  | 0.56333 | 0.53167 | 0.45112 | 0.56458 | 0.55972 | 0.5255 | 0.54717 | 0.54667 | 0.53667 | 0.555 | 0.54167 | 0.555 | 0.54333 | 0.54333 | 0.56 | 0.47667 | 0.50369 | 0.54167 | 0.47333 | 0.555 | 0.53449 | 0.53867 | 0.52 | 0.54167 | 0.57667 | 0.53833 | 0.56333
| Median |  | 0.56 | 0.53 | 0.45051 | 0.56436 | 0.56096 | 0.5245 | 0.5465 | 0.54 | 0.54 | 0.56 | 0.54 | 0.555 | 0.545 | 0.54 | 0.56 | 0.48 | 0.50343 | 0.54 | 0.485 | 0.555 | 0.52933 | 0.54 | 0.52 | 0.545 | 0.58 | 0.54 | 0.56
| Std Dev. |  | 0.005164 | 0.0040825 | 0.032218 | 0.0022605 | 0.0052093 | 0.0045056 | 0.0033116 | 0.010328 | 0.015055 | 0.0083666 | 0.0075277 | 0.0054772 | 0.008165 | 0.005164 | 0.0063246 | 0.008165 | 0.012255 | 0.017224 | 0.027325 | 0.0054772 | 0.011339 | 0.006377 | 0 | 0.0098319 | 0.005164 | 0.0075277 | 0.008165
| Rel.Std.Dev. |  | 0.9167% | 0.7679% | 7.142% | 0.4004% | 0.9307% | 0.8574% | 0.6052% | 1.889% | 2.805% | 1.507% | 1.39% | 0.9869% | 1.503% | 0.9504% | 1.129% | 1.713% | 2.433% | 3.18% | 5.773% | 0.9869% | 2.121% | 1.184% | 0% | 1.815% | 0.8955% | 1.398% | 1.449%
| PDM3 |  | 3.313% | -2.494% | -17.27% | 3.542% | 2.65% | -3.625% | 0.3484% | 0.2567% | -1.577% | 1.785% | -0.6603% | 1.785% | -0.3546% | -0.3546% | 2.702% | -12.58% | -7.625% | -0.6603% | -13.19% | 1.785% | -1.976% | -1.21% | -4.634% | -0.6603% | 5.759% | -1.272% | 3.313%
| Z-Score (Absolute) |  | 1.1 | 0.51 | 4.58 | 1.16 | 0.91 | 0.82 | 0.28 | 0.25 | 0.25 | 0.67 | 0 | 0.67 | 0.08 | 0.08 | 0.93 | 3.29 | 1.92 | 0 | 3.46 | 0.67 | 0.36 | 0.15 | 1.1 | 0 | 1.77 | 0.17 | 1.1
| Analytical results for Sb in OREAS 240b (Certified Value 1.21 ppm) | [32 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 1.76 | 1.44 | 4.61 | 1.6347 | 1.3 | 1.62 | 0.48 | 0.71 | < 5 | 1.51 | 1.36 | 1.56 | 1.52 | 101.539 | 1.27 | < 5 | < 5 | 1.34 | 1.22 | 0.9 | 7.11 | 1 | 1.6247 | 1.5 | 0.346 | 0.97 | 0.72 | 0.82 | 0.91 |  |  | 1
| 1 | 2 | 1.8 | 1.48 | 4.29 | 1.5721 | 1.36 | 1.68 | 0.42 | 0.7 | < 5 | 1.58 | 1.33 | 1.65 | 1.8 | 96.59 | 1.33 | < 5 | < 5 | 1.35 | 1.06 | 0.91 | 9.88 | 1 | 1.6958 | 1.5 | 0.356 | 0.93 | 0.72 | 0.84 | 0.9 |  |  | 32
| 1 | 3 | 1.81 | 1.46 | 2.91 | 1.4917 | 1.24 | 1.58 | 0.42 | 0.74 | < 5 | 1.8 | 1.26 | 1.6 | 1.6 | 97.608 | 1.3 | < 5 | < 5 | 1.36 | 1.08 | 0.91 | 6.74 | 1 | 1.7404 | 1.49 | 0.362 | 0.93 | 0.74 | 0.84 | 1.01 |  |  | 16
| 1 | 4 | 1.85 | 1.5 | 4.98 | 1.5216 | 1.26 | 1.58 | 0.47 | 0.75 | < 5 | 1.6 | 1.4 | 1.62 | 1.57 | 95.923 | 1.36 | < 5 | < 5 | 1.34 | 0.92 | 0.88 | 8 | 1 | 1.7603 | 1.46 | 0.329 | 0.91 | 0.75 | 0.86 | 1 |  |  | 1.2113
| 1 | 5 | 1.76 | 1.49 | 5.87 | 1.4951 | 1.31 | 1.56 | 0.41 | 0.69 | < 5 | 1.67 | 1.32 | 1.69 | 1.53 | 93.583 | 1.33 | < 5 | < 5 | 1.37 | 1.05 | 0.9 | 8.88 | 1 | 1.6264 | 1.48 | 0.344 | 0.89 | 0.74 | 0.86 | 1.05 |  |  | 115
| 1 | 6 | 1.86 | 1.49 | 5.33 | 1.476 | 1.33 | 1.63 | 0.39 | 0.7 | < 5 | 1.53 | 1.3 | 1.64 | 1.53 | 96.808 | 1.35 | < 5 | < 5 | 1.37 | 1.14 | 0.88 | 9.67 | 1 | 1.7513 | 1.47 | 0.36 | 0.84 | 0.72 | 0.84 | 1.05
| Mean |  | 1.8067 | 1.4767 | 4.665 | 1.5319 | 1.3 | 1.6083 | 0.43167 | 0.715 |  | 1.615 | 1.3283 | 1.6267 | 1.5917 | 97.0085 | 1.3233 |  |  | 1.355 | 1.0783 | 0.89667 | 8.38 | 1 | 1.6998 | 1.4833 | 0.3495 | 0.91167 | 0.73167 | 0.84333 | 0.98667
| Median |  | 1.805 | 1.485 | 4.795 | 1.5083 | 1.305 | 1.6 | 0.42 | 0.705 |  | 1.59 | 1.325 | 1.63 | 1.55 | 96.699 | 1.33 |  |  | 1.355 | 1.07 | 0.9 | 8.44 | 1 | 1.7181 | 1.485 | 0.351 | 0.92 | 0.73 | 0.84 | 1.005
| Std Dev. |  | 0.04274 | 0.022509 | 1.0215 | 0.060627 | 0.044272 | 0.044008 | 0.035449 | 0.02429 |  | 0.10672 | 0.048339 | 0.044572 | 0.10647 | 2.6082 | 0.033267 |  |  | 0.013784 | 0.10008 | 0.013663 | 1.3123 | 0 | 0.061652 | 0.01633 | 0.012422 | 0.044008 | 0.013292 | 0.015055 | 0.066533
| Rel.Std.Dev. |  | 2.366% | 1.524% | 21.9% | 3.958% | 3.406% | 2.736% | 8.212% | 3.397% |  | 6.608% | 3.639% | 2.74% | 6.689% | 2.689% | 2.514% |  |  | 1.017% | 9.281% | 1.524% | 15.66% | 0% | 3.627% | 1.101% | 3.554% | 4.827% | 1.817% | 1.785% | 6.743%
| PDM3 |  | 49.15% | 21.9% | 285.1% | 26.46% | 7.319% | 32.77% | -64.36% | -40.97% |  | 33.32% | 9.658% | 34.29% | 31.4% | 7908% | 9.245% |  |  | 11.86% | -10.98% | -25.98% | 591.8% | -17.45% | 40.33% | 22.45% | -71.15% | -24.74% | -39.6% | -30.38% | -18.55%
| Z-Score (Absolute) |  | 0.49 | 0.01 | 4.8 | 0.08 | 0.27 | 0.19 | 1.58 | 1.15 | 1.54 | 0.2 | 0.23 | 0.22 | 0.17 | 143.95 | 0.24 | 1.54 | 1.54 | 0.19 | 0.61 | 0.88 | 10.4 | NA | 0.33 | 0.01 | 1.7 | 0.86 | 1.13 | 0.96 | 0.74
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 240b (Certified Value 5.21 ppm) | [32 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 4.5 | 5.32 | 5.3 | 6.44 | 5.3217 | 6.3 | 7 | 7 | 6.2 | 5 | 4.9 | 5.4 | 5.5 | 5.2 | 4.6 | 4 | 6 | 4.4 | 4.8 | 4 | 5.2 | 2 | 5.2 | 5.1 | 4 | [6 empty] | 1
| 1 | 2 | 4.6 | 5.29 | 5.253 | 6.42 | 5.2129 | 6.8 | 7.3 | 6 | 6.2 | 5 | 4.8 | 5.3 | 5.6 | 5.1 | 4.9 | 4 | 5 | 4.4 | 4.9 | 4 | 5.1 | 2 | 5.5 | 5.1 | 4.1 | [6 empty] | 33
| 1 | 3 | 4.4 | 5.34 | 5.508 | 6.48 | 5.2544 | 6.2 | 7 | 6 | 5.9 | 5 | 4.6 | 5.2 | 5.5 | 5.1 | 4.8 | 4 | 5 | 4.1 | 4.7 | 4 | 5.2 | 2.1 | 5.3 | 5.3 | 4 | [6 empty] | 16
| 1 | 4 | 4.4 | 5.34 | 5.094 | 6.39 | 5.2191 | 6.4 | 7 | 7 | 6.2 | 5 | 4.6 | 5.4 | 5.6 | 4.9 | 5 | 4 | 5 | 3.9 | 4.9 | 4 | 5.2 | 1.9 | 5.2 | 5.1 | 3.9 | [6 empty] | 5.2114
| 1 | 5 | 4.4 | 5.3 | 5.105 | 6.4 | 5.1246 | 6.7 | 6.6 | 6 | 6 | 5 | 4.8 | 5.4 | 5.6 | 5 | 5.1 | 4 | 5 | 4.5 | 4.9 | 4 | 5.3 | 1.7 | 5.4 | 5.3 | 4.1 | [6 empty] | 116
| 1 | 6 | 4.4 | 5.28 | 5.312 | 6.43 | 5.2955 | 6.6 | 6.5 | 7 | 5.9 | 5 | 4.9 | 5.4 | 5.7 | 5.2 | 5.1 | 4 | 5 | 4.3 | 4.9 | 4 | 5.3 | 1.6 | 5.2 | 5.2 | 4.1
| Mean |  | 4.45 | 5.3117 | 5.262 | 6.4267 | 5.238 | 6.5 | 6.9 | 6.5 | 6.0667 | 5 | 4.7667 | 5.35 | 5.5833 | 5.0833 | 4.9167 | 4 | 5.1667 | 4.2667 | 4.85 | 4 | 5.2167 | 1.8833 | 5.3 | 5.1833 | 4.0333
| Median |  | 4.4 | 5.31 | 5.2765 | 6.425 | 5.2367 | 6.5 | 7 | 6.5 | 6.1 | 5 | 4.8 | 5.4 | 5.6 | 5.1 | 4.95 | 4 | 5 | 4.35 | 4.9 | 4 | 5.2 | 1.95 | 5.25 | 5.15 | 4.05
| Std Dev. |  | 0.083666 | 0.025626 | 0.15324 | 0.032042 | 0.069875 | 0.23664 | 0.29665 | 0.54772 | 0.15055 | 0 | 0.13663 | 0.083666 | 0.075277 | 0.1169 | 0.19408 | 0 | 0.40825 | 0.22509 | 0.083666 | 0 | 0.075277 | 0.19408 | 0.12649 | 0.098319 | 0.08165
| Rel.Std.Dev. |  | 1.88% | 0.4824% | 2.912% | 0.4986% | 1.334% | 3.641% | 4.299% | 8.427% | 2.482% | 0% | 2.866% | 1.564% | 1.348% | 2.3% | 3.947% | 0% | 7.902% | 5.276% | 1.725% | 0% | 1.443% | 10.31% | 2.387% | 1.897% | 2.024%
| PDM3 |  | -14.61% | 1.924% | 0.9711% | 23.32% | 0.5114% | 24.73% | 32.4% | 24.73% | 16.41% | -4.056% | -8.534% | 2.66% | 7.137% | -2.457% | -5.655% | -23.25% | -0.8582% | -18.13% | -6.935% | -23.25% | 0.1012% | -63.86% | 1.7% | -0.5384% | -22.61%
| Z-Score (Absolute) |  | 1.43 | 0.16 | 0.06 | 2.21 | 0.02 | 2.34 | 3.08 | NA | 1.54 | NA | 0.85 | 0.23 | 0.65 | 0.26 | 0.57 | NA | NA | 1.77 | 0.69 | NA | 0.02 | 6.15 | 0.13 | 0.08 | 2.2
Results from laboratories 1.10, 1.12, 1.19, 1.22 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 240b (Indicative Value 0.73 ppm) | [32 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | 0.54 | < 1 | < 1.5 | 0.8 | 3 | 1 | < 1 | < 5 | 2.4 | 0.9 | 0.5 | 0.7 | 0.8 | < 5 | 6.1852 | < 1 | 0.6 | 4.75 | < 1 | 0.486 | 0.8 | 0.51 | 2.3 | < 1 | 0.5 | 0.5 |  |  |  |  | 1
| 1 | 2 | < 1 | 0.45 | < 1 | < 1.5 | 0.8 | 11 | 1 | < 1 | < 5 | 2.4 | 0.8 | 0.6 | 0.9 | 1.1 | < 5 | < 5 | < 1 | 0.5 | 1.73 | < 1 | 0.58247 | 0.8 | 0.662 | 1.5 | < 1 | 0.5 | 0.5 |  |  |  |  | 7
| 1 | 3 | < 1 | 0.52 | < 1 | < 1.5 | 0.8 | 5 | 1 | < 1 | < 5 | 2.2 | 0.8 | 0.6 | 0.9 | 0.7 | < 5 | < 5 | < 1 | 0.6 | 1.65 | < 1 | 0.53367 | 0.8 | 0.48 | 0.8 | < 1 | 0.5 | 0.6 |  |  |  |  | 16
| 1 | 4 | < 1 | 0.56 | < 1 | < 1.5 | 0.9 | < 1 | 1 | < 1 | < 5 | 2.6 | 1.2 | 0.6 | 0.8 | 0.6 | < 5 | 5.0548 | < 1 | 0.5 | 2.97 | < 1 | 0.45519 | 1 | 0.453 | 1.2 | < 1 | 0.5 | 0.5 |  |  |  |  | 0.73208
| 1 | 5 | < 1 | 0.46 | < 1 | < 1.5 | 0.8 | < 1 | 1 | < 1 | < 5 | 2.2 | 0.9 | 0.6 | 1 | 1 | < 5 | 5.2015 | < 1 | 0.6 | 5.76 | < 1 | 0.50261 | 0.8 | 0.465 | 2.5 | < 1 | 0.6 | 0.6 |  |  |  |  | 13
| 1 | 6 | < 1 | 0.52 | < 1 | < 1.5 | 0.8 | 4 | 1 | < 1 | < 5 | 1.9 | 0.8 | 0.6 | 1 | 0.9 | < 5 | < 5 | < 1 | 0.6 | 2.48 | < 1 | 0.50279 | 0.9 | 0.544 | 1.7 | < 1 | 0.5 | 0.5
| Mean |  |  | 0.50833 |  |  | 0.81667 | 5.75 | 1 |  |  | 2.2833 | 0.9 | 0.58333 | 0.88333 | 0.85 |  | 5.4805 |  | 0.56667 | 3.2233 |  | 0.51045 | 0.85 | 0.519 | 1.6667 |  | 0.51667 | 0.53333
| Median |  |  | 0.52 |  |  | 0.8 | 4.5 | 1 |  |  | 2.3 | 0.85 | 0.6 | 0.9 | 0.85 |  | 5.2015 |  | 0.6 | 2.725 |  | 0.5027 | 0.8 | 0.495 | 1.6 |  | 0.5 | 0.5
| Std Dev. |  |  | 0.044008 |  |  | 0.040825 | 3.594 | 0 |  |  | 0.24014 | 0.15492 | 0.040825 | 0.1169 | 0.18708 |  | 0.61471 |  | 0.05164 | 1.6785 |  | 0.043584 | 0.083666 | 0.077393 | 0.64704 |  | 0.040825 | 0.05164
| Rel.Std.Dev. |  |  | 8.657% |  |  | 4.999% | 62.5% | 0% |  |  | 10.52% | 17.21% | 6.999% | 13.23% | 22.01% |  | 11.22% |  | 9.113% | 52.07% |  | 8.538% | 9.843% | 14.91% | 38.82% |  | 7.902% | 9.682%
| PDM3 |  |  | -30.56% |  |  | 11.55% | 685.4% | 36.6% |  |  | 211.9% | 22.94% | -20.32% | 20.66% | 16.11% |  | 648.6% |  | -22.6% | 340.3% |  | -30.27% | 16.11% | -29.11% | 127.7% |  | -29.43% | -27.15%
| Z-Score (Absolute) |  | 0.67 | 0.65 | 0.67 | 0 | 0.18 | 8.77 | 0.67 | 0.67 | 4.72 | 4.14 | 0.4 | 0.45 | 0.36 | 0.27 | 4.72 | 8.74 | 0.67 | 0.49 | 6.67 | 0.67 | 0.65 | 0.27 | 0.62 | 2.47 | 0.67 | 0.63 | 0.58
| Analytical results for Si in OREAS 240b (Indicative Value 0.089 wt.%) | [32 empty] | Indicative
Si | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.25 | 1.27 | [29 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | [29 empty] | 3
|  |  | 0.5g | 25g | [29 empty] | 3
| 1 | 1 | 0.14268 | 0.0655 | [29 empty] | 1
| 1 | 2 | 0.11264 | 0.06 | [29 empty] | 5
| 1 | 3 | 0.16701 | 0.0575 | [29 empty] | 16
| 1 | 4 | 0.10005 | 0.061 | [29 empty] | 0.088959
| 1 | 5 | 0.10785 | 0.0495 | [29 empty] | 11
| 1 | 6 | 0.10176 | 0.042
| Mean |  | 0.122 | 0.055917
| Median |  | 0.11025 | 0.05875
| Std Dev. |  | 0.026954 | 0.0086222
| Rel.Std.Dev. |  | 22.09% | 15.42%
| PDM3 |  | 37.14% | -37.14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sm in OREAS 240b (Certified Value 1.6 ppm) | [32 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.636 | 1.467 | 1.3205 | 2.1 | 1.75 | 1.5 | [25 empty] | 1
| 1 | 2 | 1.646 | 1.446 | 1.3846 | 1.7 | 1.81 | 1.6 | [25 empty] | 18
| 1 | 3 | 1.673 | 1.499 | 1.3386 | 1.5 | 1.74 | 1.6 | [25 empty] | 16
| 1 | 4 | 1.627 | 1.466 | 1.3246 | 1.5 | 1.8 | 1.8 | [25 empty] | 1.5952
| 1 | 5 | 1.68 | 1.514 | 1.3297 | 1.8 | 1.74 | 1.7 | [25 empty] | 117
| 1 | 6 | 1.616 | 1.581 | 1.3238 | 1.7 | 1.67 | 1.6
| Mean |  | 1.6463 | 1.4955 | 1.337 | 1.7167 | 1.7517 | 1.6333
| Median |  | 1.641 | 1.483 | 1.3272 | 1.7 | 1.745 | 1.6
| Std Dev. |  | 0.025477 | 0.048575 | 0.024184 | 0.22286 | 0.050365 | 0.10328
| Rel.Std.Dev. |  | 1.547% | 3.248% | 1.809% | 12.98% | 2.875% | 6.323%
| PDM3 |  | 3.208% | -6.247% | -16.19% | 7.617% | 9.812% | 2.393%
| Z-Score (Absolute) |  | 0.05 | 1.03 | 2.16 | 0.55 | 0.8 | 0.05
| Analytical results for Sn in OREAS 240b (Certified Value 0.75 ppm) | [32 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.8 | 0.79 | < 1 | 0.59723 | 0.69327 | 0.82 | 0.66 | 1 | 0.79 | < 10 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.7 | 44.55 | 0.5 | 0.7249 | 0.7 | 1.4 | 0.8 | 0.8 | 0.5 |  |  |  |  | 1
| 1 | 2 | 0.8 | 0.79 | < 1 | 0.61861 | 0.73917 | 0.83 | 0.7 | 0.9 | 0.82 | < 10 | 0.7 | 0.7 | 0.7 | 0.8 | 0.7 | < 10 | < 5 | 0.7 | 0.8 | 41.43 | 0.5 | 0.81916 | 0.7 | 1.3 | 0.8 | 0.8 | 0.5 |  |  |  |  | 19
| 1 | 3 | 0.8 | 0.76 | < 1 | 0.59813 | 0.78669 | 0.81 | 0.64 | 0.8 | 0.73 | < 10 | 0.7 | 0.8 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.8 | 42.18 | 0.5 | 0.80069 | 0.7 | 1.1 | 0.8 | 0.8 | 0.5 |  |  |  |  | 16
| 1 | 4 | 1 | 0.78 | < 1 | 0.57433 | 0.68058 | 0.76 | 0.67 | 0.9 | 0.76 | < 10 | 0.7 | 0.8 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.7 | 62.35 | 0.5 | 0.81201 | 0.7 | 1 | 0.8 | 0.8 | 0.5 |  |  |  |  | 0.74978
| 1 | 5 | 0.8 | 0.79 | < 1 | 0.60011 | 0.69929 | 0.8 | 0.64 | 0.9 | 0.8 | < 10 | 0.7 | 0.7 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.7 | 40.27 | 0.5 | 0.77482 | 0.7 | 0.9 | 0.8 | 0.8 | 0.5 |  |  |  |  | 118
| 1 | 6 | 0.8 | 0.76 | < 1 | 0.57143 | 0.66013 | 0.83 | 0.66 | 0.9 | 0.78 | < 10 | 0.7 | 0.8 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.7 | 0.8 | 43.74 | 0.5 | 0.77478 | 0.7 | 0.8 | 0.8 | 0.8 | 0.5
| Mean |  | 0.83333 | 0.77833 |  | 0.59331 | 0.70986 | 0.80833 | 0.66167 | 0.9 | 0.78 |  | 0.7 | 0.75 | 0.76667 | 0.71667 | 0.7 |  |  | 0.76667 | 0.75 | 45.7533 | 0.5 | 0.78439 | 0.7 | 1.0833 | 0.8 | 0.8 | 0.5
| Median |  | 0.8 | 0.785 |  | 0.59768 | 0.69628 | 0.815 | 0.66 | 0.9 | 0.785 |  | 0.7 | 0.75 | 0.8 | 0.7 | 0.7 |  |  | 0.8 | 0.75 | 42.96 | 0.5 | 0.78775 | 0.7 | 1.05 | 0.8 | 0.8 | 0.5
| Std Dev. |  | 0.08165 | 0.01472 |  | 0.017684 | 0.045781 | 0.026394 | 0.022286 | 0.063246 | 0.031623 |  | 1.2162e-16 | 0.054772 | 0.05164 | 0.040825 | 1.2162e-16 |  |  | 0.05164 | 0.054772 | 8.2762 | 0 | 0.034527 | 1.2162e-16 | 0.23166 | 1.2162e-16 | 1.2162e-16 | 0
| Rel.Std.Dev. |  | 9.798% | 1.891% |  | 2.981% | 6.449% | 3.265% | 3.368% | 7.027% | 4.054% |  | 1.737e-14% | 7.303% | 6.736% | 5.696% | 1.737e-14% |  |  | 6.736% | 7.303% | 18.09% | 0% | 4.402% | 1.737e-14% | 21.38% | 1.52e-14% | 1.52e-14% | 0%
| PDM3 |  | 11.14% | 3.808% |  | -20.87% | -5.325% | 7.809% | -11.75% | 20.03% | 4.03% |  | -6.64% | 0.02887% | 2.252% | -4.417% | -6.64% |  |  | 2.252% | 0.02887% | 6002% | -33.31% | 4.616% | -6.64% | 44.49% | 6.697% | 6.697% | -33.31%
| Z-Score (Absolute) |  | 0.67 | 0.12 | 2.7 | 1.75 | 0.57 | 0.42 | 1.06 | 1.35 | 0.13 | 42.82 | 0.67 | 0.17 | 0 | 0.51 | 0.67 | 42.82 | 17.53 | 0 | 0.17 | 455.02 | 2.7 | 0.18 | 0.67 | 3.2 | 0.34 | 0.34 | 2.7
| Analytical results for Sr in OREAS 240b (Certified Value 43.3 ppm) | [32 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 31.8 | 44.38 | 164.74 | 52.1624 | 43.8549 | 59.2 | 49 | 52.5 | 53.36 | 47 | 45 | 42.6 | 43.4 | 41.9 | 39.7 | 41 | 53.997 | 46.6 | 39.3 | 43.5 | 34.36 | 42.7521 | 33 | 43.7 | 38.6 | 40.4 | 43.8 | 34.5 |  |  |  | 1
| 1 | 2 | 29 | 44.23 | 171.35 | 51.8622 | 43.4744 | 61.3 | 51 | 50.8 | 55.31 | 48 | 46 | 42.5 | 43.6 | 41.4 | 42.3 | 40 | 51.5503 | 47.4 | 40 | 44.1 | 37.21 | 45.9513 | 32 | 42.9 | 39.4 | 39.5 | 43.6 | 34.2 |  |  |  | 20
| 1 | 3 | 29.5 | 44.86 | 152.31 | 49.9987 | 43.99 | 57.8 | 51 | 51.4 | 53.25 | 47 | 43 | 41 | 43.1 | 41.1 | 42 | 43 | 54.5724 | 49 | 37.4 | 42.1 | 36.06 | 45.3654 | 33 | 43.5 | 39.3 | 42 | 44.4 | 34.6 |  |  |  | 16
| 1 | 4 | 31.9 | 45.25 | 167.59 | 50.5756 | 44.1081 | 58.1 | 50 | 56.4 | 55.04 | 48 | 44 | 42.1 | 44.2 | 40.2 | 43.4 | 44 | 51.3507 | 48.8 | 34.4 | 42.5 | 36.86 | 41.2641 | 32 | 44 | 38.3 | 41.1 | 43.1 | 34.9 |  |  |  | 43.2605
| 1 | 5 | 31.9 | 44.45 | 183.77 | 49.3413 | 42.6012 | 59.1 | 49 | 52.4 | 53.52 | 47 | 44 | 42.5 | 45.5 | 41.1 | 44.5 | 43 | 50.4037 | 49.8 | 39.8 | 43.1 | 36.55 | 42.7614 | 31 | 44.8 | 37.2 | 40.9 | 44.6 | 33.9 |  |  |  | 119
| 1 | 6 | 30.8 | 43.71 | 163.59 | 49.7003 | 43.9268 | 61.2 | 52 | 54.6 | 53.34 | 47 | 44 | 42.7 | 44.6 | 41.6 | 44 | 40 | 49.2052 | 46.1 | 37.8 | 43.4 | 37.77 | 42.4824 | 30 | 45.4 | 36.7 | 40.7 | 42.9 | 34.4
| Mean |  | 30.8167 | 44.48 | 167.225 | 50.6067 | 43.6592 | 59.45 | 50.3333 | 53.0167 | 53.97 | 47.3333 | 44.3333 | 42.2333 | 44.0667 | 41.2167 | 42.65 | 41.8333 | 51.8466 | 47.95 | 38.1167 | 43.1167 | 36.4683 | 43.4295 | 31.8333 | 44.05 | 38.25 | 40.7667 | 43.7333 | 34.4167
| Median |  | 31.3 | 44.415 | 166.165 | 50.2871 | 43.8908 | 59.15 | 50.5 | 52.45 | 53.44 | 47 | 44 | 42.5 | 43.9 | 41.25 | 42.85 | 42 | 51.4505 | 48.1 | 38.55 | 43.25 | 36.705 | 42.7567 | 32 | 43.85 | 38.45 | 40.8 | 43.7 | 34.45
| Std Dev. |  | 1.2922 | 0.52983 | 10.3221 | 1.1653 | 0.56111 | 1.4977 | 1.2111 | 2.1037 | 0.94132 | 0.5164 | 1.0328 | 0.6377 | 0.88919 | 0.58452 | 1.7352 | 1.7224 | 2.0711 | 1.4694 | 2.1075 | 0.72226 | 1.1846 | 1.822 | 1.169 | 0.9094 | 1.1005 | 0.82381 | 0.6802 | 0.34303
| Rel.Std.Dev. |  | 4.193% | 1.191% | 6.173% | 2.303% | 1.285% | 2.519% | 2.406% | 3.968% | 1.744% | 1.091% | 2.33% | 1.51% | 2.018% | 1.418% | 4.069% | 4.117% | 3.995% | 3.064% | 5.529% | 1.675% | 3.248% | 4.195% | 3.672% | 2.064% | 2.877% | 2.021% | 1.555% | 0.9967%
| PDM3 |  | -28.76% | 2.819% | 286.6% | 16.98% | 0.9216% | 37.42% | 16.35% | 22.55% | 24.76% | 9.415% | 2.48% | -2.374% | 1.863% | -4.725% | -1.411% | -3.299% | 19.85% | 10.84% | -11.89% | -0.3325% | -15.7% | 0.3905% | -26.41% | 1.825% | -11.58% | -5.765% | 1.093% | -20.44%
| Z-Score (Absolute) |  | 2.2 | 0.13 | 21.11 | 1.18 | 0.01 | 2.69 | 1.13 | 1.59 | 1.76 | 0.62 | 0.11 | 0.25 | 0.06 | 0.42 | 0.18 | 0.32 | 1.39 | 0.73 | 0.95 | 0.1 | 1.24 | 0.05 | 2.03 | 0.06 | 0.93 | 0.5 | 0.01 | 1.59
| Analytical results for Ta in OREAS 240b (Certified Value < 0.01 ppm) | [32 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.24 | 0.06 | < 0.01 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.45 | 0.13 | < 0.01 | < 0.01 | [11 empty] | 21
| 1 | 3 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.23 | 0.05 | < 0.01 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | 0.012779 | < 0.05 | 0.02 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.25 | 0.15 | < 0.01 | < 0.01 | [11 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.01 | 0.010029 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | < 0.01 | 0.28 | 0.06 | < 0.01 | < 0.01 | [11 empty] | 120
| 1 | 6 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.02 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.19 | 0.08 | < 0.01 | < 0.01
| Mean |  |  |  | 0.011404 |  | 0.026667 |  |  | 5 |  | 0.01 | 0.01 |  |  |  |  | 0.01 | 0.27333 | 0.088333
| Median |  |  |  | 0.011404 |  | 0.03 |  |  | 5 |  | 0.01 | 0.01 |  |  |  |  | 0.01 | 0.245 | 0.07
| Std Dev. |  |  |  | 0.0019444 |  | 0.005164 |  |  | 0 |  | 0 | 0 |  |  |  |  | 0 | 0.09136 | 0.041673
| Rel.Std.Dev. |  |  |  | 17.05% |  | 19.36% |  |  | 0% |  | 0% | 0% |  |  |  |  | 0% | 33.42% | 47.18%
PDM3
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.56 | 0.4 | 0.49 | 0.4 | 0.67 | 268.78 | 1.75 | 0.58 | 0.4 | 0.67 | 0.67 | 0.4 | 12.54 | 0.45 | 13.8 | 3.82 | 0.67 | 0.67
| Analytical results for Tb in OREAS 240b (Certified Value 0.29 ppm) | [32 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 3
| 1 | 1 | 0.29 | 0.304 | 0.285 | 0.28013 | 0.3 | 0.28 | 0.25 | 0.29 | [23 empty] | 1
| 1 | 2 | 0.3 | 0.283 | 0.277 | 0.28063 | 0.3 | 0.3 | 0.25 | 0.3 | [23 empty] | 22
| 1 | 3 | 0.28 | 0.296 | 0.291 | 0.28216 | 0.3 | 0.29 | 0.25 | 0.31 | [23 empty] | 16
| 1 | 4 | 0.28 | 0.303 | 0.281 | 0.28319 | 0.3 | 0.28 | 0.25 | 0.3 | [23 empty] | 0.29002
| 1 | 5 | 0.28 | 0.297 | 0.297 | 0.27622 | 0.3 | 0.29 | 0.25 | 0.31 | [23 empty] | 121
| 1 | 6 | 0.28 | 0.292 | 0.306 | 0.28648 | 0.3 | 0.28 | 0.25 | 0.3
| Mean |  | 0.285 | 0.29583 | 0.2895 | 0.28147 | 0.3 | 0.28667 | 0.25 | 0.30167
| Median |  | 0.28 | 0.2965 | 0.288 | 0.28139 | 0.3 | 0.285 | 0.25 | 0.3
| Std Dev. |  | 0.0083666 | 0.0077309 | 0.010766 | 0.003423 | 0 | 0.008165 | 0 | 0.0075277
| Rel.Std.Dev. |  | 2.936% | 2.613% | 3.719% | 1.216% | 0% | 2.848% | 0% | 2.495%
| PDM3 |  | -1.732% | 2.004% | -0.1802% | -2.949% | 3.44% | -1.157% | -13.8% | 4.015%
| Z-Score (Absolute) |  | 0.22 | 1.19 | 0.37 | 0.67 | NA | 0 | 4.76 | 1.95
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 240b (Certified Value 0.14 ppm) | [32 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 0.1 | 0.14 | 0.15149 | 0.16 | 0.14 | < 0.05 | 0.1 | < 5 | 0.1 | 0.13 | 0.14 | 0.14 | 0.15 | < 5 | 0.14 | 0.14 | 0.1 | 0.18192 | 0.14 | 0.077 | 0.13 | 0.13 | 0.11 | 0.15 | [7 empty] | 1
| 1 | 2 | 0.18 | 0.15 | 0.163 | 0.13 | 0.14 | 0.07 | 0.2 | < 5 | 0.3 | 0.13 | 0.15 | 0.15 | 0.12 | < 5 | 0.14 | 0.13 | 0.1 | 0.10926 | 0.14 | 0.103 | 0.14 | 0.14 | 0.13 | 0.1 | [7 empty] | 23
| 1 | 3 | 0.09 | 0.14 | 0.15443 | 0.17 | 0.12 | < 0.05 | 0.1 | < 5 | 0.1 | 0.12 | 0.15 | 0.13 | 0.12 | < 5 | 0.14 | 0.13 | 0.1 | 0.16683 | 0.15 | 0.154 | 0.03 | 0.14 | 0.12 | 0.17 | [7 empty] | 16
| 1 | 4 | 0.13 | 0.14 | 0.14555 | 0.09 | 0.13 | 0.08 | 0.1 | < 5 | 0.2 | 0.15 | 0.16 | 0.15 | 0.15 | < 5 | 0.15 | 0.11 | 0.1 | 0.15407 | 0.14 | < 0.03 | 0.1 | 0.15 | 0.13 | 0.07 | [7 empty] | 0.13814
| 1 | 5 | 0.1 | 0.14 | 0.17081 | 0.12 | 0.14 | 0.08 | 0.1 | < 5 | 0.1 | 0.14 | 0.14 | 0.13 | 0.14 | < 5 | 0.15 | 0.14 | 0.1 | 0.12792 | 0.15 | 0.103 | 0.09 | 0.14 | 0.12 | 0.14 | [7 empty] | 122
| 1 | 6 | 0.14 | 0.14 | 0.16498 | 0.13 | 0.13 | 0.07 | 0.1 | < 5 | 0.1 | 0.16 | 0.15 | 0.15 | 0.13 | < 5 | 0.14 | 0.13 | 0.1 | 0.12476 | 0.14 | 0.077 | < 0.01 | 0.13 | 0.12 | 0.12
| Mean |  | 0.12333 | 0.14167 | 0.15838 | 0.13333 | 0.13333 | 0.075 | 0.11667 |  | 0.15 | 0.13833 | 0.14833 | 0.14167 | 0.135 |  | 0.14333 | 0.13 | 0.1 | 0.14413 | 0.14333 | 0.1028 | 0.098 | 0.13833 | 0.12167 | 0.125
| Median |  | 0.115 | 0.14 | 0.15872 | 0.13 | 0.135 | 0.075 | 0.1 |  | 0.1 | 0.135 | 0.15 | 0.145 | 0.135 |  | 0.14 | 0.13 | 0.1 | 0.141 | 0.14 | 0.103 | 0.1 | 0.14 | 0.12 | 0.13
| Std Dev. |  | 0.033862 | 0.0040825 | 0.0094534 | 0.028752 | 0.008165 | 0.0057735 | 0.040825 |  | 0.083666 | 0.01472 | 0.0075277 | 0.0098319 | 0.013784 |  | 0.005164 | 0.010954 | 1.5202e-17 | 0.027914 | 0.005164 | 0.031436 | 0.043243 | 0.0075277 | 0.0075277 | 0.036194
| Rel.Std.Dev. |  | 27.46% | 2.882% | 5.969% | 21.56% | 6.124% | 7.698% | 34.99% |  | 55.78% | 10.64% | 5.075% | 6.94% | 10.21% |  | 3.603% | 8.427% | 1.52e-14% | 19.37% | 3.603% | 30.58% | 44.13% | 5.442% | 6.187% | 28.96%
| PDM3 |  | -10.72% | 2.556% | 14.65% | -3.477% | -3.477% | -45.71% | -15.54% |  | 8.589% | 0.1431% | 7.382% | 2.556% | -2.27% |  | 3.763% | -5.89% | -27.61% | 4.337% | 3.763% | -25.58% | -29.06% | 0.1431% | -11.92% | -9.509%
| Z-Score (Absolute) |  | 1.21 | 0.27 | 1.62 | 0.4 | 0.4 | 6.47 | NA | 191.1 | NA | 0 | 0.81 | 0.27 | 0.27 | 191.1 | 0.4 | 0.67 | NA | 0.47 | 0.4 | 4.06 | 4.52 | 0 | 1.35 | 1.08
Results from laboratories 1.11, 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 240b (Certified Value 1.14 ppm) | [32 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.8 | 1.208 | 0.988 | 0.80778 | 1.1979 | 1.2 | 1.23 | 1.32 | 1.1 | 1.2 | 1 | 1.1 | 1.1 | < 100 | 1.1 | 1.08 | 1 | 1.2 | 1.1 | 1.2 | 3 | 1.1 | 1.1 | 1.1 | [7 empty] | 1
| 1 | 2 | 0.8 | 1.165 | 1.031 | 0.80778 | 1.2113 | 1.2 | 1.32 | 1.32 | 1.16 | 1.3 | 1.1 | 1.2 | 1.2 | < 100 | 1.1 | 1.1 | 1 | 1.2 | 1 | 1.1 | 2.4 | 1 | 1.2 | 1.1 | [7 empty] | 24
| 1 | 3 | 0.8 | 1.174 | 0.986 | 0.80044 | 1.2122 | 1.2 | 1.21 | 1.3 | 1.09 | 1.3 | 1 | 1.1 | 1.1 | < 100 | 1.1 | 1.1 | 0.9 | 1.2 | 1 | 1.1 | 2.2 | 1.1 | 1.2 | 1.1 | [7 empty] | 16
| 1 | 4 | 0.8 | 1.209 | 0.985 | 0.77654 | 1.2098 | 1.2 | 1.24 | 1.34 | 1.12 | 1.3 | 1.1 | 1.2 | 1.2 | < 100 | 1.1 | 1.12 | 0.9 | 1.2 | 1.2 | 1.1 | 1.9 | 1 | 1.1 | 1.1 | [7 empty] | 1.1379
| 1 | 5 | 0.8 | 1.188 | 1.04 | 0.7808 | 1.1883 | 1.2 | 1.21 | 1.34 | 1.14 | 1.3 | 1.1 | 1.1 | 1.1 | < 100 | 1.1 | 1.13 | 1 | 1.2 | 1.1 | 1.1 | 1.8 | 1.1 | 1.2 | 1.1 | [7 empty] | 123
| 1 | 6 | 0.8 | 1.189 | 1.06 | 0.77582 | 1.2135 | 1.2 | 1.18 | 1.33 | 1.12 | 1.3 | 1 | 1.1 | 1.1 | < 100 | 1.1 | 1.1 | 1 | 1.2 | 1.1 | 1.1 | 1.7 | 1.1 | 1.1 | 1.1
| Mean |  | 0.8 | 1.1888 | 1.015 | 0.79153 | 1.2055 | 1.2 | 1.2317 | 1.325 | 1.1217 | 1.2833 | 1.05 | 1.1333 | 1.1333 |  | 1.1 | 1.105 | 0.96667 | 1.2 | 1.0833 | 1.1167 | 2.1667 | 1.0667 | 1.15 | 1.1
| Median |  | 0.8 | 1.1885 | 1.0095 | 0.79062 | 1.2105 | 1.2 | 1.22 | 1.325 | 1.12 | 1.3 | 1.05 | 1.1 | 1.1 |  | 1.1 | 1.1 | 1 | 1.2 | 1.1 | 1.1 | 2.05 | 1.1 | 1.15 | 1.1
| Std Dev. |  | 1.2162e-16 | 0.01768 | 0.03279 | 0.015456 | 0.010129 | 0 | 0.047924 | 0.015166 | 0.025626 | 0.040825 | 0.054772 | 0.05164 | 0.05164 |  | 2.4324e-16 | 0.017607 | 0.05164 | 0 | 0.075277 | 0.040825 | 0.48442 | 0.05164 | 0.054772 | 2.4324e-16
| Rel.Std.Dev. |  | 1.52e-14% | 1.487% | 3.231% | 1.953% | 0.8402% | 0% | 3.891% | 1.145% | 2.285% | 3.181% | 5.216% | 4.556% | 4.556% |  | 2.211e-14% | 1.593% | 5.342% | 0% | 6.949% | 3.656% | 22.36% | 4.841% | 4.763% | 2.211e-14%
| PDM3 |  | -29.7% | 4.475% | -10.8% | -30.44% | 5.938% | 5.456% | 8.239% | 16.44% | -1.428% | 12.78% | -7.726% | -0.4027% | -0.4027% |  | -3.332% | -2.893% | -15.05% | 5.456% | -4.797% | -1.867% | 90.41% | -6.261% | 1.062% | -3.332%
| Z-Score (Absolute) |  | 3.05 | 0.57 | 1.05 | 3.12 | 0.73 | 0.67 | 0.97 | 1.84 | 0.05 | 1.45 | 0.72 | 0.05 | 0.05 | 454.55 | 0.26 | 0.21 | 1.5 | 0.67 | 0.41 | 0.1 | 9.67 | 0.57 | 0.21 | 0.26
| Analytical results for Ti in OREAS 240b (Certified Value 0.282 wt.%) | [32 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.29 | 0.2495 | 0.34759 | 0.37787 | 0.29185 | 0.34 | 0.3273 | 0.386 | 0.3378 | 0.28 | 0.234 | 0.263 | 0.276 | 0.262 | < 0.01 | 0.239 | 0.21 | 0.39858 | 0.3239 | 0.226 | 0.25 | 0.18861 | 0.27587 | 0.1025 | 0.263 | 0.24 | 0.27 | 0.264 | 0.194 |  |  | 1
| 1 | 2 | 0.28 | 0.2527 | 0.32214 | 0.37653 | 0.28721 | 0.35 | 0.3489 | 0.374 | 0.3483 | 0.28 | 0.242 | 0.258 | 0.254 | 0.258 | < 0.01 | 0.252 | 0.2 | 0.3854 | 0.3177 | 0.224 | 0.25 | 0.23498 | 0.29593 | 0.096 | 0.265 | 0.249 | 0.28 | 0.26 | 0.197 |  |  | 25
| 1 | 3 | 0.27 | 0.2562 | 0.34559 | 0.37866 | 0.29287 | 0.34 | 0.3495 | 0.385 | 0.3364 | 0.29 | 0.211 | 0.267 | 0.259 | 0.258 | < 0.01 | 0.257 | 0.23 | 0.4008 | 0.3212 | 0.196 | 0.24 | 0.21728 | 0.29921 | 0.103 | 0.263 | 0.249 | 0.27 | 0.271 | 0.195 |  |  | 16
| 1 | 4 | 0.27 | 0.2543 | 0.33236 | 0.37897 | 0.28605 | 0.34 | 0.3441 | 0.389 | 0.3386 | 0.29 | 0.214 | 0.262 | 0.285 | 0.259 | < 0.01 | 0.256 | 0.23 | 0.37587 | 0.3368 | 0.198 | 0.25 | 0.23366 | 0.26248 | 0.098 | 0.261 | 0.234 | 0.27 | 0.259 | 0.191 |  |  | 0.28216
| 1 | 5 | 0.26 | 0.2552 | 0.34815 | 0.37658 | 0.27938 | 0.35 | 0.3285 | 0.379 | 0.3307 | 0.28 | 0.22 | 0.261 | 0.278 | 0.252 | < 0.01 | 0.263 | 0.23 | 0.37423 | 0.3424 | 0.236 | 0.24 | 0.22234 | 0.27043 | 0.0895 | 0.268 | 0.213 | 0.28 | 0.275 | 0.186 |  |  | 124
| 1 | 6 | 0.27 | 0.2473 | 0.31559 | 0.38023 | 0.29162 | 0.34 | 0.3579 | 0.38 | 0.3314 | 0.29 | 0.223 | 0.259 | 0.283 | 0.262 | < 0.01 | 0.261 | 0.21 | 0.36661 | 0.3364 | 0.226 | 0.26 | 0.24123 | 0.27359 | 0.0945 | 0.271 | 0.204 | 0.27 | 0.261 | 0.182
| Mean |  | 0.27333 | 0.25253 | 0.33524 | 0.37814 | 0.28816 | 0.34333 | 0.3427 | 0.38217 | 0.3372 | 0.285 | 0.224 | 0.26167 | 0.2725 | 0.2585 |  | 0.25467 | 0.21833 | 0.38358 | 0.32973 | 0.21767 | 0.24833 | 0.22301 | 0.27959 | 0.09725 | 0.26517 | 0.2315 | 0.27333 | 0.265 | 0.19083
| Median |  | 0.27 | 0.2535 | 0.33898 | 0.37826 | 0.28942 | 0.34 | 0.3465 | 0.3825 | 0.3371 | 0.285 | 0.2215 | 0.2615 | 0.277 | 0.2585 |  | 0.2565 | 0.22 | 0.38064 | 0.33015 | 0.225 | 0.25 | 0.228 | 0.27473 | 0.097 | 0.264 | 0.237 | 0.27 | 0.2625 | 0.1925
| Std Dev. |  | 0.010328 | 0.0034726 | 0.014087 | 0.001444 | 0.0051023 | 0.005164 | 0.0123 | 0.0054924 | 0.006357 | 0.0054772 | 0.011916 | 0.0032042 | 0.012911 | 0.0036742 |  | 0.0085946 | 0.013292 | 0.013855 | 0.010065 | 0.016561 | 0.0075277 | 0.019002 | 0.014688 | 0.0051064 | 0.0037103 | 0.018918 | 0.005164 | 0.0065422 | 0.0057764
| Rel.Std.Dev. |  | 3.779% | 1.375% | 4.202% | 0.3819% | 1.771% | 1.504% | 3.589% | 1.437% | 1.885% | 1.922% | 5.32% | 1.225% | 4.738% | 1.421% |  | 3.375% | 6.088% | 3.612% | 3.052% | 7.608% | 3.031% | 8.52% | 5.253% | 5.251% | 1.399% | 8.172% | 1.889% | 2.469% | 3.027%
| PDM3 |  | -3.129% | -10.5% | 18.81% | 34.02% | 2.127% | 21.68% | 21.46% | 35.44% | 19.51% | 1.006% | -20.61% | -7.264% | -3.424% | -8.386% |  | -9.744% | -22.62% | 35.94% | 16.86% | -22.86% | -11.99% | -20.96% | -0.9131% | -65.53% | -6.023% | -17.95% | -3.129% | -6.082% | -32.37%
| Z-Score (Absolute) |  | 0.13 | 0.21 | 1.15 | 1.85 | 0.38 | 1.28 | 1.27 | 1.92 | 1.18 | 0.32 | 0.67 | 0.06 | 0.12 | 0.11 | 4.26 | 0.17 | 0.77 | 1.94 | 1.06 | 0.78 | 0.28 | 0.69 | 0.24 | 2.75 | 0 | 0.55 | 0.13 | 0 | 1.22
| Analytical results for Tl in OREAS 240b (Certified Value 0.18 ppm) | [32 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.18 | 0.19 | 0.18225 | 0.18 | 0.18 | 0.17 | < 10 | 0.2 | 0.18 | 0.17 | 0.2 | 0.17 | < 5 | 0.18 | 0.17 | 0.18 | 5.32 | 0.15 | 0.18522 | 0.19 | 0.189 | 0.14 | 0.16 | 0.17 | 0.17 | [6 empty] | 1
| 1 | 2 | 0.18 | 0.18 | 0.18793 | 0.18 | 0.19 | 0.18 | < 10 | 0.2 | 0.19 | 0.18 | 0.19 | 0.18 | < 5 | 0.16 | 0.18 | 0.18 | 9.34 | 0.15 | 0.20652 | 0.18 | 0.189 | 0.12 | 0.17 | 0.18 | 0.17 | [6 empty] | 26
| 1 | 3 | 0.18 | 0.19 | 0.18164 | 0.18 | 0.19 | 0.18 | < 10 | 0.19 | 0.19 | 0.17 | 0.18 | 0.17 | < 5 | 0.18 | 0.17 | 0.18 | 12.35 | 0.15 | 0.18271 | 0.18 | 0.192 | 0.12 | 0.17 | 0.18 | 0.16 | [6 empty] | 16
| 1 | 4 | 0.18 | 0.19 | 0.20583 | 0.17 | 0.18 | 0.17 | < 10 | 0.19 | 0.2 | 0.18 | 0.16 | 0.18 | < 5 | 0.16 | 0.15 | 0.18 | 8.09 | 0.15 | 0.19674 | 0.19 | 0.181 | 0.1 | 0.17 | 0.18 | 0.17 | [6 empty] | 0.18027
| 1 | 5 | 0.17 | 0.19 | 0.18835 | 0.17 | 0.18 | 0.17 | < 10 | 0.21 | 0.17 | 0.18 | 0.19 | 0.18 | < 5 | 0.17 | 0.17 | 0.18 | 3.82 | 0.15 | 0.18335 | 0.18 | 0.197 | 0.1 | 0.17 | 0.18 | 0.18 | [6 empty] | 125
| 1 | 6 | 0.18 | 0.19 | 0.18932 | 0.18 | 0.19 | 0.17 | < 10 | 0.2 | 0.18 | 0.18 | 0.16 | 0.17 | < 5 | 0.18 | 0.17 | 0.19 | 3.99 | 0.15 | 0.18828 | 0.18 | 0.196 | 0.1 | 0.16 | 0.18 | 0.18
| Mean |  | 0.17833 | 0.18833 | 0.18922 | 0.17667 | 0.185 | 0.17333 |  | 0.19833 | 0.185 | 0.17667 | 0.18 | 0.175 |  | 0.17167 | 0.16833 | 0.18167 | 7.1517 | 0.15 | 0.19047 | 0.18333 | 0.19067 | 0.11333 | 0.16667 | 0.17833 | 0.17167
| Median |  | 0.18 | 0.19 | 0.18814 | 0.18 | 0.185 | 0.17 |  | 0.2 | 0.185 | 0.18 | 0.185 | 0.175 |  | 0.175 | 0.17 | 0.18 | 6.705 | 0.15 | 0.18675 | 0.18 | 0.1905 | 0.11 | 0.17 | 0.18 | 0.17
| Std Dev. |  | 0.0040825 | 0.0040825 | 0.0087679 | 0.005164 | 0.0054772 | 0.005164 |  | 0.0075277 | 0.010488 | 0.005164 | 0.016733 | 0.0054772 |  | 0.0098319 | 0.0098319 | 0.0040825 | 3.3806 | 0 | 0.0093828 | 0.005164 | 0.0058195 | 0.01633 | 0.005164 | 0.0040825 | 0.0075277
| Rel.Std.Dev. |  | 2.289% | 2.168% | 4.634% | 2.923% | 2.961% | 2.979% |  | 3.795% | 5.669% | 2.923% | 9.296% | 3.13% |  | 5.727% | 5.841% | 2.247% | 47.27% | 0% | 4.926% | 2.817% | 3.052% | 14.41% | 3.098% | 2.289% | 4.385%
| PDM3 |  | -1.073% | 4.474% | 4.967% | -1.998% | 2.625% | -3.847% |  | 10.02% | 2.625% | -1.998% | -0.1489% | -2.923% |  | -4.772% | -6.621% | 0.7756% | 3867% | -16.79% | 5.659% | 1.7% | 5.768% | -37.13% | -7.545% | -1.073% | -4.772%
| Z-Score (Absolute) |  | 0.13 | 0.67 | 0.75 | 0.27 | 0.4 | 0.54 | 390.02 | 1.48 | 0.4 | 0.27 | 0 | 0.4 | 187.73 | 0.67 | 0.94 | 0.13 | 564.13 | 2.43 | 0.85 | 0.27 | 0.86 | 5.39 | 1.08 | 0.13 | 0.67
| Analytical results for Tm in OREAS 240b (Certified Value 0.14 ppm) | [32 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | [26 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [26 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [26 empty] | 2
| 1 | 1 | 0.152 | 0.146 | 0.13359 | 0.2 | 0.1 | [26 empty] | 1
| 1 | 2 | 0.158 | 0.14 | 0.12766 | 0.2 | 0.1 | [26 empty] | 27
| 1 | 3 | 0.159 | 0.149 | 0.13311 | 0.2 | 0.1 | [26 empty] | 16
| 1 | 4 | 0.161 | 0.143 | 0.12904 | 0.1 | 0.1 | [26 empty] | 0.14058
| 1 | 5 | 0.163 | 0.155 | 0.12962 | 0.2 | 0.1 | [26 empty] | 126
| 1 | 6 | 0.151 | 0.157 | 0.13031 | 0.1 | 0.1
| Mean |  | 0.15733 | 0.14833 | 0.13056 | 0.16667 | 0.1
| Median |  | 0.1585 | 0.1475 | 0.12997 | 0.2 | 0.1
| Std Dev. |  | 0.0048442 | 0.0066833 | 0.0023374 | 0.05164 | 1.5202e-17
| Rel.Std.Dev. |  | 3.079% | 4.506% | 1.79% | 30.98% | 1.52e-14%
| PDM3 |  | 11.92% | 5.517% | -7.129% | 18.56% | -28.87%
| Z-Score (Absolute) |  | 0.34 | 0 | 0.67 | 0.7 | 1.83
| Analytical results for U in OREAS 240b (Certified Value 0.28 ppm) | [32 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.3 | 0.294 | 0.278 | 0.2059 | 0.32923 | 0.3 | 0.32 | 0.37 | 0.3 | 0.3 | 0.27 | 0.27 | 0.28 | 1491.182 | 0.27 | < 10 | 0.26 | 0.2 | 0.31 | 0.2 | 0.29 | 0.276 | 0.26 | 0.28 | 0.29 | 0.25 | [5 empty] | 1
| 1 | 2 | 0.29 | 0.297 | 0.282 | 0.20205 | 0.34112 | 0.3 | 0.37 | 0.36 | 0.32 | 0.3 | 0.27 | 0.27 | 0.27 | 1424.208 | 0.28 | < 10 | 0.28 | 0.2 | 0.3 | 0.2 | 0.28 | 0.271 | 0.26 | 0.29 | 0.3 | 0.26 | [5 empty] | 28
| 1 | 3 | 0.29 | 0.294 | 0.283 | 0.21074 | 0.31933 | 0.3 | 0.37 | 0.36 | 0.3 | 0.3 | 0.26 | 0.27 | 0.28 | 1341.37 | 0.27 | < 10 | 0.29 | 0.2 | 0.29 | 0.2 | 0.27 | 0.274 | 0.27 | 0.28 | 0.31 | 0.25 | [5 empty] | 16
| 1 | 4 | 0.29 | 0.31 | 0.278 | 0.20564 | 0.31359 | 0.3 | 0.34 | 0.38 | 0.31 | 0.3 | 0.27 | 0.27 | 0.29 | 1355.562 | 0.28 | < 10 | 0.3 | 0.2 | 0.3 | 0.3 | 0.28 | 0.26 | 0.26 | 0.28 | 0.29 | 0.26 | [5 empty] | 0.28384
| 1 | 5 | 0.29 | 0.295 | 0.292 | 0.20396 | 0.3095 | 0.3 | 0.35 | 0.36 | 0.31 | 0.3 | 0.27 | 0.29 | 0.28 | 1339.374 | 0.29 | < 10 | 0.3 | 0.2 | 0.29 | 0.2 | 0.27 | 0.285 | 0.26 | 0.28 | 0.3 | 0.25 | [5 empty] | 127
| 1 | 6 | 0.29 | 0.294 | 0.299 | 0.20272 | 0.31067 | 0.3 | 0.33 | 0.37 | 0.3 | 0.3 | 0.26 | 0.28 | 0.28 | 1346.2 | 0.29 | < 10 | 0.26 | 0.3 | 0.3 | 0.2 | 0.28 | 0.306 | 0.25 | 0.27 | 0.3 | 0.25
| Mean |  | 0.29167 | 0.29733 | 0.28533 | 0.20517 | 0.32057 | 0.3 | 0.34667 | 0.36667 | 0.30667 | 0.3 | 0.26667 | 0.275 | 0.28 | 1382.9827 | 0.28 |  | 0.28167 | 0.21667 | 0.29833 | 0.21667 | 0.27833 | 0.27867 | 0.26 | 0.28 | 0.29833 | 0.25333
| Median |  | 0.29 | 0.2945 | 0.2825 | 0.2048 | 0.31646 | 0.3 | 0.345 | 0.365 | 0.305 | 0.3 | 0.27 | 0.27 | 0.28 | 1350.881 | 0.28 |  | 0.285 | 0.2 | 0.3 | 0.2 | 0.28 | 0.275 | 0.26 | 0.28 | 0.3 | 0.25
| Std Dev. |  | 0.0040825 | 0.006314 | 0.0084301 | 0.0031285 | 0.012395 | 0 | 0.020656 | 0.008165 | 0.008165 | 0 | 0.005164 | 0.0083666 | 0.0063246 | 61.8786 | 0.0089443 |  | 0.018348 | 0.040825 | 0.0075277 | 0.040825 | 0.0075277 | 0.015642 | 0.0063246 | 0.0063246 | 0.0075277 | 0.005164
| Rel.Std.Dev. |  | 1.4% | 2.124% | 2.954% | 1.525% | 3.867% | 0% | 5.958% | 2.227% | 2.662% | 0% | 1.936% | 3.042% | 2.259% | 4.474% | 3.194% |  | 6.514% | 18.84% | 2.523% | 18.84% | 2.705% | 5.613% | 2.433% | 2.259% | 2.523% | 2.038%
| PDM3 |  | 2.758% | 4.754% | 0.5265% | -27.72% | 12.94% | 5.694% | 22.13% | 29.18% | 8.042% | 5.694% | -6.05% | -3.114% | -1.353% | 487142% | -1.353% |  | -0.7653% | -23.67% | 5.107% | -23.67% | -1.94% | -1.822% | -8.399% | -1.353% | 5.107% | -10.75%
| Z-Score (Absolute) |  | 0.37 | 0.63 | 0.08 | 3.56 | 1.69 | NA | 2.87 | 3.78 | 1.05 | NA | 0.77 | 0.39 | 0.16 | 62856.15 | 0.16 | 214.41 | 0.08 | NA | 0.67 | NA | 0.23 | 0.22 | 1.07 | 0.16 | 0.67 | 1.37
Results from laboratories 1.07, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 240b (Certified Value 116 ppm) | [32 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 122 | 119.8 | 116.92 | 124.9822 | 126.0232 | 126 | 142 | 154 | 144 | 119 | 117 | 111 | 113 | 109 | < 10 | 105 | 99 | 155.7908 | 132 | 106 | 112 | 91.8 | 120.014 | 80 | 113 | 106 | 110 | 121 | 87.1 |  |  | 1
| 1 | 2 | 122 | 119.4 | 110.02 | 123.1577 | 124.4553 | 131 | 151 | 155 | 151 | 121 | 118 | 108 | 108 | 108 | < 10 | 110 | 95 | 151.6032 | 132 | 104 | 111 | 104.06 | 128.938 | 75 | 113 | 106 | 112 | 119 | 86.8 |  |  | 29
| 1 | 3 | 117 | 120.4 | 115.21 | 122.4295 | 125.8097 | 127 | 141 | 152 | 144 | 119 | 109 | 111 | 108 | 108 | < 10 | 112 | 104 | 157.3175 | 133 | 99 | 111 | 99.24 | 127.4108 | 80 | 113 | 108 | 111 | 123 | 87.1 |  |  | 16
| 1 | 4 | 117 | 120.4 | 112.51 | 120.5427 | 124.5632 | 124 | 139 | 157 | 146 | 120 | 110 | 109 | 115 | 108 | < 10 | 110 | 106 | 151.0311 | 135 | 92 | 113 | 104.64 | 116.3248 | 75 | 114 | 102 | 111 | 119 | 87.9 |  |  | 115.6936
| 1 | 5 | 112 | 120.4 | 116.31 | 122.6614 | 121.6074 | 127 | 144 | 153 | 143 | 121 | 117 | 110 | 114 | 106 | < 10 | 113 | 105 | 147.1807 | 135 | 109 | 113 | 101.53 | 119.8208 | 75 | 113 | 96.6 | 112 | 123 | 85.6 |  |  | 128
| 1 | 6 | 116 | 119.8 | 108.16 | 120.5623 | 127.0873 | 130 | 135 | 154 | 141 | 119 | 113 | 109 | 116 | 109 | < 10 | 112 | 98 | 144.7509 | 133 | 101 | 116 | 106.93 | 119.1216 | 70 | 115 | 94.4 | 110 | 121 | 86.9
| Mean |  | 117.6667 | 120.0333 | 113.1883 | 122.3893 | 124.9243 | 127.5 | 142 | 154.1667 | 144.8333 | 119.8333 | 114 | 109.6667 | 112.3333 | 108 |  | 110.3333 | 101.1667 | 151.279 | 133.3333 | 101.8333 | 112.6667 | 101.3667 | 121.9383 | 75.8333 | 113.5 | 102.1667 | 111 | 121 | 86.9
| Median |  | 117 | 120.1 | 113.86 | 122.5455 | 125.1864 | 127 | 141.5 | 154 | 144 | 119.5 | 115 | 109.5 | 113.5 | 108 |  | 111 | 101.5 | 151.3171 | 133 | 102.5 | 112.5 | 102.795 | 119.9174 | 75 | 113 | 104 | 111 | 121 | 87
| Std Dev. |  | 3.8297 | 0.4274 | 3.5648 | 1.6815 | 1.8985 | 2.5884 | 5.3666 | 1.7224 | 3.4303 | 0.98319 | 3.8987 | 1.2111 | 3.5024 | 1.0954 |  | 2.8752 | 4.446 | 4.8249 | 1.3663 | 5.9805 | 1.8619 | 5.385 | 5.0314 | 3.7639 | 0.83666 | 5.5633 | 0.89443 | 1.7889 | 0.74565
| Rel.Std.Dev. |  | 3.255% | 0.3561% | 3.149% | 1.374% | 1.52% | 2.03% | 3.779% | 1.117% | 2.368% | 0.8205% | 3.42% | 1.104% | 3.118% | 1.014% |  | 2.606% | 4.395% | 3.189% | 1.025% | 5.873% | 1.653% | 5.312% | 4.126% | 4.963% | 0.7371% | 5.445% | 0.8058% | 1.478% | 0.8581%
| PDM3 |  | 1.705% | 3.751% | -2.165% | 5.787% | 7.979% | 10.2% | 22.74% | 33.25% | 25.19% | 3.578% | -1.464% | -5.209% | -2.904% | -6.65% |  | -4.633% | -12.56% | 30.76% | 15.25% | -11.98% | -2.616% | -12.38% | 5.398% | -34.45% | -1.896% | -11.69% | -4.057% | 4.587% | -24.89%
| Z-Score (Absolute) |  | 0.32 | 0.5 | 0.02 | 0.67 | 0.87 | 1.06 | 2.16 | 3.08 | 2.38 | 0.48 | 0.04 | 0.29 | 0.09 | 0.42 | 8.23 | 0.24 | 0.94 | 2.87 | 1.5 | 0.88 | 0.06 | 0.92 | 0.64 | 2.86 | 0 | 0.86 | 0.19 | 0.57 | 2.02
| Analytical results for W in OREAS 240b (Certified Value 33.7 ppm) | [32 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 37.8 | 30.15 | 42.13 | 40.9924 | 27.9 | 39 | 25.1 | 24.92 | 29 | 35.9 | 36 | 34.4 | 34.5 | < 100 | 33.3 | 26 | 34.1 | 27 | 36 | 28 | 36.6 | 39.2 | 35.7 | 30.73 | 39.5 | [6 empty] | 1
| 1 | 2 | 39.2 | 29.85 | 41.68 | 40.6283 | 29.7 | 38 | 25.8 | 24.96 | 29 | 36.1 | 36.7 | 35.1 | 35.9 | < 100 | 34.9 | 26 | 34 | 26.2 | 37 | 29 | 37.4 | 40.2 | 36.4 | 30.59 | 39.7 | [6 empty] | 30
| 1 | 3 | 38.6 | 29.78 | 41.15 | 41.9191 | 27 | 40 | 25.8 | 25.69 | 29 | 36.2 | 36.3 | 34.9 | 35.6 | < 100 | 33.9 | 27 | 35.5 | 24.7 | 36 | 29 | 36.5 | 40.5 | 37.7 | 31.6 | 39.7 | [6 empty] | 16
| 1 | 4 | 39.2 | 30.3 | 41.32 | 41.2906 | 26.2 | 39 | 24.4 | 25.27 | 29 | 35.7 | 36 | 36.7 | 36 | < 100 | 34.6 | 26 | 35.1 | 22.7 | 36 | 29 | 36.5 | 40.4 | 36.6 | 29.46 | 40.4 | [6 empty] | 33.7297
| 1 | 5 | 39.1 | 29.84 | 43.14 | 39.7814 | 26.6 | 40 | 23.7 | 23.95 | 29 | 37.7 | 35.8 | 37.7 | 36.5 | < 100 | 34.7 | 26 | 35.6 | 24.6 | 37 | 28 | 36 | 39.9 | 37.1 | 31.43 | 39.6 | [6 empty] | 129
| 1 | 6 | 39 | 29.71 | 40.17 | 42.0222 | 27.4 | 39 | 23.8 | 23.46 | 29 | 36 | 35 | 36.4 | 34.7 | < 100 | 34.5 | 27 | 34.4 | 27.2 | 36 | 28 | 36.2 | 39.3 | 37 | 30.06 | 40.4
| Mean |  | 38.8167 | 29.9383 | 41.5983 | 41.1057 | 27.4667 | 39.1667 | 24.7667 | 24.7083 | 29 | 36.2667 | 35.9667 | 35.8667 | 35.5333 |  | 34.3167 | 26.3333 | 34.7833 | 25.4 | 36.3333 | 28.5 | 36.5333 | 39.9167 | 36.75 | 30.645 | 39.8833
| Median |  | 39.05 | 29.845 | 41.5 | 41.1415 | 27.2 | 39 | 24.75 | 24.94 | 29 | 36.05 | 36 | 35.75 | 35.75 |  | 34.55 | 26 | 34.75 | 25.45 | 36 | 28.5 | 36.5 | 40.05 | 36.8 | 30.66 | 39.7
| Std Dev. |  | 0.54559 | 0.2325 | 0.99847 | 0.84027 | 1.2453 | 0.75277 | 0.94375 | 0.83913 | 0 | 0.72296 | 0.56804 | 1.266 | 0.78145 |  | 0.60139 | 0.5164 | 0.70828 | 1.7239 | 0.5164 | 0.54772 | 0.48028 | 0.55648 | 0.68337 | 0.81055 | 0.40702
| Rel.Std.Dev. |  | 1.406% | 0.7766% | 2.4% | 2.044% | 4.534% | 1.922% | 3.811% | 3.396% | 0% | 1.993% | 1.579% | 3.53% | 2.199% |  | 1.752% | 1.961% | 2.036% | 6.787% | 1.421% | 1.922% | 1.315% | 1.394% | 1.86% | 2.645% | 1.021%
| PDM3 |  | 15.08% | -11.24% | 23.33% | 21.87% | -18.57% | 16.12% | -26.57% | -26.75% | -14.02% | 7.522% | 6.632% | 6.336% | 5.347% |  | 1.74% | -21.93% | 3.124% | -24.7% | 7.719% | -15.5% | 8.312% | 18.34% | 8.954% | -9.145% | 18.24%
| Z-Score (Absolute) |  | 0.49 | 0.99 | 0.95 | 0.87 | 1.4 | 0.55 | 1.85 | 1.86 | 1.14 | 0.07 | 0.02 | 0 | 0.06 | 2.35 | 0.26 | 1.59 | 0.18 | 1.74 | 0.08 | 1.23 | 0.11 | 0.67 | 0.15 | 0.87 | 0.67
| Analytical results for Y in OREAS 240b (Certified Value 10 ppm) | [32 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 10.23 | 10.66 | 10.44 | 9.5856 | 10.2685 | 10.6 | 10.4 | 12.8 | 10.26 | 9 | 10.5 | 9.72 | 9.83 | 9.88 | 9.29 | 9 | 10.84 | 9.6 | 8 | 10.65 | 9.17 | 9.5 | 10.39 | 8.76 | [7 empty] | 1
| 1 | 2 | 10.28 | 10.67 | 10.06 | 9.6243 | 10.1658 | 10.7 | 10.4 | 12.4 | 10.33 | 9 | 10.6 | 9.61 | 10 | 9.7 | 9.86 | 8 | 10.79 | 9.8 | 8 | 10.5 | 9.18 | 9.2 | 10.33 | 8.66 | [7 empty] | 31
| 1 | 3 | 10.08 | 10.66 | 10.31 | 9.691 | 10.1852 | 10.2 | 10 | 12.7 | 9.95 | 9 | 10.1 | 9.43 | 10.1 | 9.71 | 9.65 | 9 | 10.86 | 9.63 | 8.2 | 10.5 | 9.41 | 9.9 | 10.59 | 8.78 | [7 empty] | 16
| 1 | 4 | 10.04 | 10.67 | 10.16 | 9.5569 | 10.1215 | 9.68 | 10.1 | 13.3 | 10.19 | 9 | 10.3 | 9.97 | 10.4 | 9.56 | 10.05 | 9 | 10.77 | 9.97 | 7.9 | 10.7 | 9.2 | 9.6 | 10.23 | 8.87 | [7 empty] | 10.0103
| 1 | 5 | 9.89 | 10.66 | 10.38 | 9.5108 | 9.9879 | 10.6 | 10.1 | 12.8 | 10.13 | 9 | 10.7 | 9.7 | 10.05 | 9.63 | 10.2 | 9 | 10.82 | 9.7 | 7.7 | 10.7 | 8.79 | 9.5 | 10.63 | 8.58 | [7 empty] | 130
| 1 | 6 | 9.99 | 10.65 | 9.94 | 9.5806 | 10.299 | 10.9 | 9.76 | 12.7 | 9.82 | 9 | 10.4 | 9.8 | 10 | 9.7 | 10.15 | 9 | 10.88 | 9.88 | 7.5 | 10.8 | 8.66 | 9.4 | 10.24 | 8.7
| Mean |  | 10.085 | 10.6617 | 10.215 | 9.5915 | 10.1713 | 10.4467 | 10.1267 | 12.7833 | 10.1133 | 9 | 10.4333 | 9.705 | 10.0633 | 9.6967 | 9.8667 | 8.8333 | 10.8267 | 9.7633 | 7.8833 | 10.6417 | 9.0683 | 9.5167 | 10.4017 | 8.725
| Median |  | 10.06 | 10.66 | 10.235 | 9.5831 | 10.1755 | 10.6 | 10.1 | 12.75 | 10.16 | 9 | 10.45 | 9.71 | 10.025 | 9.7 | 9.955 | 9 | 10.83 | 9.75 | 7.95 | 10.675 | 9.175 | 9.5 | 10.36 | 8.73
| Std Dev. |  | 0.14707 | 0.0075277 | 0.1947 | 0.061383 | 0.11145 | 0.43939 | 0.24549 | 0.29269 | 0.19356 | 0 | 0.21602 | 0.18119 | 0.18833 | 0.10671 | 0.34772 | 0.40825 | 0.041793 | 0.14569 | 0.24833 | 0.12007 | 0.28323 | 0.23166 | 0.17233 | 0.10114
| Rel.Std.Dev. |  | 1.458% | 0.07061% | 1.906% | 0.64% | 1.096% | 4.206% | 2.424% | 2.29% | 1.914% | 0% | 2.071% | 1.867% | 1.871% | 1.1% | 3.524% | 4.622% | 0.386% | 1.492% | 3.15% | 1.128% | 3.123% | 2.434% | 1.657% | 1.159%
| PDM3 |  | 0.7465% | 6.507% | 2.045% | -4.183% | 1.609% | 4.359% | 1.163% | 27.7% | 1.03% | -10.09% | 4.226% | -3.05% | 0.53% | -3.133% | -1.435% | -11.76% | 8.156% | -2.467% | -21.25% | 6.307% | -9.41% | -4.931% | 3.91% | -12.84%
| Z-Score (Absolute) |  | 0.03 | 1.02 | 0.21 | 0.92 | 0.13 | 0.63 | 0.05 | 4.88 | 0.03 | NA | 0.61 | 0.72 | 0.07 | 0.73 | 0.42 | NA | 1.32 | 0.61 | 4.03 | 0.99 | 1.87 | 1.06 | 0.55 | 2.5
| Analytical results for Yb in OREAS 240b (Certified Value 0.91 ppm) | [32 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 2
| 1 | 1 | 0.9 | 0.996 | 0.816 | 0.86571 | 1 | 1.11 | 0.75 | 0.9 | [23 empty] | 1
| 1 | 2 | 0.9 | 0.955 | 0.768 | 0.84106 | 1.1 | 1.15 | 0.75 | 0.8 | [23 empty] | 32
| 1 | 3 | 0.9 | 0.993 | 0.828 | 0.81398 | 1 | 1.13 | 0.75 | 0.9 | [23 empty] | 16
| 1 | 4 | 0.9 | 0.999 | 0.784 | 0.80982 | 1 | 1.09 | 0.75 | 0.9 | [23 empty] | 0.90634
| 1 | 5 | 0.9 | 0.992 | 0.807 | 0.8401 | 1 | 1.12 | 0.7 | 0.8 | [23 empty] | 131
| 1 | 6 | 0.9 | 0.967 | 1.052 | 0.82808 | 1 | 1.06 | 0.65 | 0.9
| Mean |  | 0.9 | 0.98367 | 0.8425 | 0.83312 | 1.0167 | 1.11 | 0.725 | 0.86667
| Median |  | 0.9 | 0.9925 | 0.8115 | 0.83409 | 1 | 1.115 | 0.75 | 0.9
| Std Dev. |  | 0 | 0.018129 | 0.10491 | 0.020535 | 0.040825 | 0.031623 | 0.041833 | 0.05164
| Rel.Std.Dev. |  | 0% | 1.843% | 12.45% | 2.465% | 4.016% | 2.849% | 5.77% | 5.958%
| PDM3 |  | -0.6996% | 8.532% | -7.044% | -8.078% | 12.17% | 22.47% | -20.01% | -4.377%
| Z-Score (Absolute) |  | 0.15 | 0.9 | 0.37 | 0.45 | 1.19 | 2.03 | 1.42 | 0.15
| Analytical results for Zn in OREAS 240b (Certified Value 158 ppm) | [32 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 161 | 153.7 | 144.52 | 152.85 | 164.5866 | 157 | 162 | 162 | 160 | 150 | 172 | 166 | 155 | 160 | < 100 | 170 | 148 | 151.7705 | 161 | 158.6 | 153 | 155.28 | 153.7 | 150 | 177.5341 | 165 | 145.68 | 154 | 161 | 152 | 159 | 1
| 1 | 2 | 163 | 156.9 | 142.06 | 149.94 | 163.619 | 163 | 179 | 159 | 162 | 149 | 173 | 163 | 160 | 164 | < 100 | 174 | 142 | 152.8049 | 160 | 154.5 | 156 | 149.87 | 161.43 | 150 | 175.0408 | 165 | 145.489 | 156 | 164 | 151 | 157 | 33
| 1 | 3 | 163 | 157 | 142.99 | 152.17 | 166.0716 | 161 | 158 | 161 | 155 | 148 | 172 | 167 | 158 | 161 | < 100 | 174 | 149 | 152.9663 | 160 | 154.5 | 154 | 150.52 | 161.65 | 155 | 176.6254 | 166 | 146.105 | 150 | 170 | 155 | 159 | 16
| 1 | 4 | 164 | 157.9 | 141.19 | 150.71 | 163.205 | 156 | 160 | 165 | 163 | 148 | 173 | 163 | 159 | 159 | < 100 | 172 | 147 | 150.5567 | 161 | 138.2 | 154 | 152.78 | 154.71 | 150 | 166.6577 | 166 | 140.858 | 154 | 165 | 151 | 162 | 158.0831
| 1 | 5 | 161 | 157.9 | 141.08 | 150.92 | 162.6345 | 161 | 163 | 161 | 156 | 150 | 183 | 166 | 161 | 161 | < 100 | 175 | 145 | 148.8173 | 160 | 158.3 | 156 | 150.15 | 162.18 | 150 | 180.2078 | 166 | 137.24 | 160 | 163 | 153 | 157 | 132
| 1 | 6 | 163 | 154.3 | 140.99 | 154.31 | 164.3271 | 161 | 155 | 160 | 156 | 147 | 174 | 165 | 158 | 159 | < 100 | 173 | 146 | 147.4784 | 161 | 155.1 | 154 | 158.04 | 158.53 | 150 | 175.7588 | 166 | 146.071 | 154 | 162 | 152 | 162
| Mean |  | 162.5 | 156.2833 | 142.1383 | 151.8167 | 164.074 | 159.8333 | 162.8333 | 161.3333 | 158.6667 | 148.6667 | 174.5 | 165 | 158.5 | 160.6667 |  | 173 | 146.1667 | 150.7323 | 160.5 | 153.2 | 154.5 | 152.7733 | 158.7 | 150.8333 | 175.3041 | 165.6667 | 143.5738 | 154.6667 | 164.1667 | 152.3333 | 159.3333
| Median |  | 163 | 156.95 | 141.625 | 151.545 | 163.9731 | 161 | 161 | 161 | 158 | 148.5 | 173 | 165.5 | 158.5 | 160.5 |  | 173.5 | 146.5 | 151.1636 | 160.5 | 154.8 | 154 | 151.65 | 159.98 | 150 | 176.1921 | 166 | 145.5845 | 154 | 163.5 | 152 | 159
| Std Dev. |  | 1.2247 | 1.8291 | 1.3958 | 1.61 | 1.2126 | 2.7142 | 8.4242 | 2.0656 | 3.4448 | 1.2111 | 4.2308 | 1.6733 | 2.0736 | 1.8619 |  | 1.7889 | 2.4833 | 2.2204 | 0.54772 | 7.5784 | 1.2247 | 3.2988 | 3.7206 | 2.0412 | 4.6 | 0.5164 | 3.6943 | 3.266 | 3.1885 | 1.5055 | 2.2509
| Rel.Std.Dev. |  | 0.7537% | 1.17% | 0.982% | 1.06% | 0.739% | 1.698% | 5.173% | 1.28% | 2.171% | 0.8146% | 2.425% | 1.014% | 1.308% | 1.159% |  | 1.034% | 1.699% | 1.473% | 0.3413% | 4.947% | 0.7927% | 2.159% | 2.344% | 1.353% | 2.624% | 0.3117% | 2.573% | 2.112% | 1.942% | 0.9883% | 1.413%
| PDM3 |  | 2.794% | -1.138% | -10.09% | -3.964% | 3.79% | 1.107% | 3.005% | 2.056% | 0.3692% | -5.957% | 10.39% | 4.376% | 0.2637% | 1.634% |  | 9.436% | -7.538% | -4.65% | 1.529% | -3.089% | -2.267% | -3.359% | 0.3903% | -4.586% | 10.89% | 4.797% | -9.178% | -2.161% | 3.848% | -3.637% | 0.7909%
| Z-Score (Absolute) |  | 0.47 | 0.29 | 2.03 | 0.84 | 0.66 | 0.14 | 0.51 | 0.33 | 0 | 1.23 | 1.94 | 0.78 | 0.02 | 0.25 | 13.32 | 1.76 | 1.53 | 0.97 | 0.22 | 0.67 | 0.51 | 0.72 | 0 | 0.96 | 2.04 | 0.86 | 1.85 | 0.49 | 0.67 | 0.78 | 0.08
| Analytical results for Zr in OREAS 240b (Certified Value 14.5 ppm) | [32 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 13.6 | 15.08 | 21.89 | 15.2367 | 14.3 | 13.1 | 17.9 | 18.4 | 13 | 12.6 | 13.8 | 15.7 | 14.2 | 13.7 | 12 | 14.7 | 13.1 | 16.722 | 8 | 15.6 | 11.9 | 13.8 | 15.6 | 11.5 | [7 empty] | 1
| 1 | 2 | 13.2 | 15.1 | 20.22 | 15.0899 | 14.7 | 12.7 | 18.3 | 19.1 | 13 | 13.2 | 13.5 | 15.4 | 14 | 14.5 | 12 | 13.8 | 13.1 | 18.144 | 8 | 15.5 | 12.5 | 13.5 | 15.4 | 11.2 | [7 empty] | 34
| 1 | 3 | 12.3 | 15.63 | 21.54 | 15.2963 | 12.9 | 14.5 | 19.1 | 18.1 | 13 | 12.3 | 13.5 | 15.6 | 13.8 | 14.1 | 13 | 14.4 | 12.1 | 18.144 | 8 | 15.5 | 12.7 | 14.3 | 15.9 | 11.1 | [7 empty] | 16
| 1 | 4 | 13 | 15.38 | 23.03 | 14.897 | 11.4 | 14.8 | 17.1 | 18.8 | 13 | 11.4 | 14.3 | 16 | 13.7 | 14.8 | 13 | 15.2 | 12.5 | 15.993 | 8 | 15.5 | 12.1 | 13.7 | 15.4 | 11.1 | [7 empty] | 14.4571
| 1 | 5 | 12.6 | 15.25 | 21.27 | 14.5907 | 13 | 13.2 | 18 | 18.2 | 13 | 12.4 | 13.5 | 16.1 | 14.1 | 15 | 13 | 14 | 12.6 | 16.623 | 7 | 15.6 | 11.9 | 13.9 | 16.3 | 10.8 | [7 empty] | 133
| 1 | 6 | 12.6 | 15.11 | 19.96 | 15.2125 | 12.2 | 13.3 | 17.4 | 18 | 13 | 12 | 14.1 | 15.9 | 14 | 14.8 | 12 | 15.1 | 12.9 | 16.758 | 7 | 15.9 | 11.4 | 13.6 | 15.2 | 10.9
| Mean |  | 12.8833 | 15.2583 | 21.3183 | 15.0539 | 13.0833 | 13.6 | 17.9667 | 18.4333 | 13 | 12.3167 | 13.7833 | 15.7833 | 13.9667 | 14.4833 | 12.5 | 14.5333 | 12.7167 | 17.064 | 7.6667 | 15.6 | 12.0833 | 13.8 | 15.6333 | 11.1
| Median |  | 12.8 | 15.18 | 21.405 | 15.1512 | 12.95 | 13.25 | 17.95 | 18.3 | 13 | 12.35 | 13.65 | 15.8 | 14 | 14.65 | 12.5 | 14.55 | 12.75 | 16.74 | 8 | 15.55 | 12 | 13.75 | 15.5 | 11.1
| Std Dev. |  | 0.47504 | 0.21535 | 1.1278 | 0.26753 | 1.2449 | 0.8438 | 0.70333 | 0.43205 | 0 | 0.60139 | 0.34881 | 0.26394 | 0.18619 | 0.49565 | 0.54772 | 0.57155 | 0.392 | 0.88147 | 0.5164 | 0.15492 | 0.46655 | 0.28284 | 0.40332 | 0.24495
| Rel.Std.Dev. |  | 3.687% | 1.411% | 5.29% | 1.777% | 9.515% | 6.204% | 3.915% | 2.344% | 0% | 4.883% | 2.531% | 1.672% | 1.333% | 3.422% | 4.382% | 3.933% | 3.083% | 5.166% | 6.736% | 0.9931% | 3.861% | 2.05% | 2.58% | 2.207%
| PDM3 |  | -10.89% | 5.542% | 47.46% | 4.127% | -9.503% | -5.929% | 24.28% | 27.5% | -10.08% | -14.81% | -4.661% | 9.173% | -3.393% | 0.1812% | -13.54% | 0.527% | -12.04% | 18.03% | -46.97% | 7.905% | -16.42% | -4.545% | 8.136% | -23.22%
| Z-Score (Absolute) |  | 0.83 | 0.4 | 3.55 | 0.3 | 0.73 | 0.46 | 1.81 | 2.05 | NA | 1.12 | 0.36 | 0.67 | 0.27 | 0 | NA | 0.03 | 0.92 | 1.34 | NA | 0.58 | 1.24 | 0.35 | 0.6 | 1.75
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C550: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C568: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1059: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1077: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1095: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1132: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1204: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 240b (Indicative Value 12.32 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.32 | 1
| 1 | 2 | 12.32 | 9
| Mean |  | 12.32 | 16
| Median |  | 12.32 | 12.32
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 240b (Indicative Value 8.16 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 8.16 | 1
| 1 | 2 | 8.16 | 10
| Mean |  | 8.16 | 16
| Median |  | 8.16 | 8.16
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 240b (Indicative Value 10.83 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.82 | 1
| 1 | 2 | 10.83 | 11
| Mean |  | 10.825 | 16
| Median |  | 10.825 | 10.825
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.06532%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 240b (Indicative Value 0.853 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.852 | 1
| 1 | 2 | 0.853 | 12
| Mean |  | 0.8525 | 16
| Median |  | 0.8525 | 0.8525
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.08295%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 240b (Indicative Value 5.86 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.87 | 1
| 1 | 2 | 5.85 | 9
| Mean |  | 5.86 | 16
| Median |  | 5.86 | 5.86
| Std Dev. |  | 0.014142 | 15
| Rel.Std.Dev. |  | 0.2413%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 240b (Indicative Value 0.169 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.169 | 1
| 1 | 2 | 0.168 | 10
| Mean |  | 0.1685 | 16
| Median |  | 0.1685 | 0.1685
| Std Dev. |  | 0.00070711 | 16
| Rel.Std.Dev. |  | 0.4196%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 240b (Indicative Value 2.5 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.51 | 1
| 1 | 2 | 2.49 | 11
| Mean |  | 2.5 | 16
| Median |  | 2.5 | 2.5
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.5657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 240b (Indicative Value 0.095 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.095 | 1
| 1 | 2 | 0.094 | 12
| Mean |  | 0.0945 | 16
| Median |  | 0.0945 | 0.0945
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 240b (Indicative Value 0.457 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.4525 | 1
| 1 | 2 | 0.4605 | 16
| Mean |  | 0.4565 | 16
| Median |  | 0.4565 | 0.45651
| Std Dev. |  | 0.0056569 | 15
| Rel.Std.Dev. |  | 1.239%
| PDM3 |  | -0.002848%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 240b (Indicative Value 54.61 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 54.58 | 1
| 1 | 2 | 54.64 | 10
| Mean |  | 54.61 | 16
| Median |  | 54.61 | 54.61
| Std Dev. |  | 0.042426 | 16
| Rel.Std.Dev. |  | 0.07769%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 240b (Indicative Value 0.957 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.957 | 1
| 1 | 2 | 0.957 | 11
| Mean |  | 0.957 | 16
| Median |  | 0.957 | 0.957
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 240b (Indicative Value 3.36 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.36 | 1
| 1 | 2 | 3.35 | 14
| Mean |  | 3.355 | 16
| Median |  | 3.355 | 3.355
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.2108%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 240b (Indicative Value 0.19 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.19 | 1
| 1 | 2 | 0.19 | 16
| Mean |  | 0.19 | 16
| Median |  | 0.19 | 0.19
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 240b (Indicative Value 0.515 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.51 | 1
| 1 | 2 | 0.52 | 16
| Mean |  | 0.515 | 16
| Median |  | 0.515 | 0.515
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 1.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 240b (Indicative Value 1.3 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.3 | 18
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 240b (Indicative Value 94 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 93.2 | 1
| 1 | 2 | 94.2 | 19
| Mean |  | 93.7 | 16
| Median |  | 93.7 | 93.7
| Std Dev. |  | 0.70711 | 25
| Rel.Std.Dev. |  | 0.7546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 240b (Indicative Value 314 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 314 | 1
| 1 | 2 | 313 | 20
| Mean |  | 313.5 | 16
| Median |  | 313.5 | 313.5
| Std Dev. |  | 0.70711 | 26
| Rel.Std.Dev. |  | 0.2256%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 240b (Indicative Value 0.7 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.8 | 21
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0.14142 | 27
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 240b (Indicative Value 0.07 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.08 | 22
| Mean |  | 0.07 | 16
| Median |  | 0.07 | 0.07
| Std Dev. |  | 0.014142 | 28
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 240b (Indicative Value 0.95 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.9 | 1
| 1 | 2 | 1 | 23
| Mean |  | 0.95 | 16
| Median |  | 0.95 | 0.95
| Std Dev. |  | 0.070711 | 29
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 240b (Indicative Value 14.4 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.3 | 1
| 1 | 2 | 14.4 | 24
| Mean |  | 14.35 | 16
| Median |  | 14.35 | 14.35
| Std Dev. |  | 0.070711 | 30
| Rel.Std.Dev. |  | 0.4928%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 240b (Indicative Value 41.2 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 41.1 | 1
| 1 | 2 | 41.3 | 25
| Mean |  | 41.2 | 16
| Median |  | 41.2 | 41.2
| Std Dev. |  | 0.14142 | 31
| Rel.Std.Dev. |  | 0.3433%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 240b (Indicative Value 165 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 165 | 26
| Mean |  | 164.5 | 16
| Median |  | 164.5 | 164.5
| Std Dev. |  | 0.70711 | 32
| Rel.Std.Dev. |  | 0.4299%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 240b (Indicative Value 1.32 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.34 | 1
| 1 | 2 | 1.3 | 27
| Mean |  | 1.32 | 16
| Median |  | 1.32 | 1.32
| Std Dev. |  | 0.028284 | 33
| Rel.Std.Dev. |  | 2.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 240b (Indicative Value 169 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 166 | 1
| 1 | 2 | 172 | 28
| Mean |  | 169 | 16
| Median |  | 169 | 169
| Std Dev. |  | 4.2426 | 34
| Rel.Std.Dev. |  | 2.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 240b (Indicative Value 3.45 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.42 | 1
| 1 | 2 | 3.47 | 29
| Mean |  | 3.445 | 16
| Median |  | 3.445 | 3.445
| Std Dev. |  | 0.035355 | 35
| Rel.Std.Dev. |  | 1.026%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 240b (Indicative Value 2.2 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.19 | 1
| 1 | 2 | 2.2 | 30
| Mean |  | 2.195 | 16
| Median |  | 2.195 | 2.195
| Std Dev. |  | 0.0070711 | 36
| Rel.Std.Dev. |  | 0.3221%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 240b (Indicative Value 0.88 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.88 | 1
| 1 | 2 | 0.87 | 31
| Mean |  | 0.875 | 16
| Median |  | 0.875 | 0.875
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 240b (Indicative Value 14.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.2 | 1
| 1 | 2 | 14.4 | 32
| Mean |  | 14.3 | 16
| Median |  | 14.3 | 14.3
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 0.989%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 240b (Indicative Value 3.11 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.11 | 1
| 1 | 2 | 3.1 | 33
| Mean |  | 3.105 | 16
| Median |  | 3.105 | 3.105
| Std Dev. |  | 0.0070711 | 39
| Rel.Std.Dev. |  | 0.2277%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 240b (Indicative Value 1.45 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.5 | 34
| Mean |  | 1.45 | 16
| Median |  | 1.45 | 1.45
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 240b (Indicative Value 2.02 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.01 | 1
| 1 | 2 | 2.03 | 18
| Mean |  | 2.02 | 16
| Median |  | 2.02 | 2.02
| Std Dev. |  | 0.014142 | 24
| Rel.Std.Dev. |  | 0.7001%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 240b (Indicative Value 0.78 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.79 | 19
| Mean |  | 0.775 | 16
| Median |  | 0.775 | 0.775
| Std Dev. |  | 0.021213 | 25
| Rel.Std.Dev. |  | 2.737%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 240b (Indicative Value 0.075 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.1 | 20
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.035355 | 26
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 240b (Indicative Value 6.69 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.68 | 1
| 1 | 2 | 6.7 | 21
| Mean |  | 6.69 | 16
| Median |  | 6.69 | 6.69
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 0.2114%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 240b (Indicative Value 0.32 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.32 | 22
| Mean |  | 0.32 | 16
| Median |  | 0.32 | 0.32
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 240b (Indicative Value 0.148 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.147 | 1
| 1 | 2 | 0.149 | 23
| Mean |  | 0.148 | 16
| Median |  | 0.148 | 0.148
| Std Dev. |  | 0.0014142 | 29
| Rel.Std.Dev. |  | 0.9555%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 240b (Indicative Value 2.9 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.2 | 1
| 1 | 2 | 2.6 | 24
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.42426 | 30
| Rel.Std.Dev. |  | 14.63%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 240b (Indicative Value 3.54 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.54 | 1
| 1 | 2 | 3.54 | 7
| Mean |  | 3.54 | 16
| Median |  | 3.54 | 3.54
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 240b (Indicative Value 9.03 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.02 | 1
| 1 | 2 | 9.04 | 26
| Mean |  | 9.03 | 16
| Median |  | 9.03 | 9.03
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 0.1566%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 240b (Indicative Value 98 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 96 | 27
| Mean |  | 98 | 16
| Median |  | 98 | 98
| Std Dev. |  | 2.8284 | 33
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 240b (Indicative Value 40.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 41 | 28
| Mean |  | 40.5 | 16
| Median |  | 40.5 | 40.5
| Std Dev. |  | 0.70711 | 34
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 240b (Indicative Value 1.99 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 1.98 | 29
| Mean |  | 1.99 | 16
| Median |  | 1.99 | 1.99
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 0.7107%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 240b (Indicative Value 20 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.8 | 1
| 1 | 2 | 20.2 | 30
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0.28284 | 36
| Rel.Std.Dev. |  | 1.414%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 240b (Indicative Value 0.02 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.02 | 31
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.02
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 240b (Indicative Value 2.2 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2.3 | 32
| Mean |  | 2.2 | 16
| Median |  | 2.2 | 2.2
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 6.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 240b (Indicative Value 34.1 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.8 | 1
| 1 | 2 | 34.4 | 33
| Mean |  | 34.1 | 16
| Median |  | 34.1 | 34.1
| Std Dev. |  | 0.42426 | 39
| Rel.Std.Dev. |  | 1.244%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 240b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 7
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 40
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 240b (Indicative Value 2.54 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.61 | 1
| 1 | 2 | 2.46 | 18
| Mean |  | 2.535 | 16
| Median |  | 2.535 | 2.535
| Std Dev. |  | 0.10607 | 24
| Rel.Std.Dev. |  | 4.184%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 240b (Indicative Value 1.2 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 19
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 240b (Indicative Value 188 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 189 | 1
| 1 | 2 | 186 | 20
| Mean |  | 187.5 | 16
| Median |  | 187.5 | 187.5
| Std Dev. |  | 2.1213 | 26
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 240b (Indicative Value 0.23 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.23 | 1
| 1 | 2 | 0.22 | 21
| Mean |  | 0.225 | 16
| Median |  | 0.225 | 0.225
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 240b (Indicative Value 0.54 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.54 | 1
| 1 | 2 | 0.54 | 22
| Mean |  | 0.54 | 16
| Median |  | 0.54 | 0.54
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 240b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 29
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 240b (Indicative Value 1.46 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.52 | 1
| 1 | 2 | 1.4 | 24
| Mean |  | 1.46 | 16
| Median |  | 1.46 | 1.46
| Std Dev. |  | 0.084853 | 30
| Rel.Std.Dev. |  | 5.812%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 240b (Indicative Value 0.566 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.567 | 1
| 1 | 2 | 0.565 | 25
| Mean |  | 0.566 | 16
| Median |  | 0.566 | 0.566
| Std Dev. |  | 0.0014142 | 31
| Rel.Std.Dev. |  | 0.2499%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 240b (Indicative Value 0.15 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | < 0.2 | 26
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 32
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 240b (Indicative Value 0.33 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.34 | 27
| Mean |  | 0.33 | 16
| Median |  | 0.33 | 0.33
| Std Dev. |  | 0.014142 | 33
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 240b (Indicative Value 0.39 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.39 | 28
| Mean |  | 0.385 | 16
| Median |  | 0.385 | 0.385
| Std Dev. |  | 0.0070711 | 34
| Rel.Std.Dev. |  | 1.837%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 240b (Indicative Value 241 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 241 | 1
| 1 | 2 | 240 | 29
| Mean |  | 240.5 | 16
| Median |  | 240.5 | 240.5
| Std Dev. |  | 0.70711 | 35
| Rel.Std.Dev. |  | 0.294%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 240b (Indicative Value 46.3 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 46.5 | 30
| Mean |  | 46.25 | 16
| Median |  | 46.25 | 46.25
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 0.7644%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 240b (Indicative Value 19.3 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.1 | 1
| 1 | 2 | 19.4 | 31
| Mean |  | 19.25 | 16
| Median |  | 19.25 | 19.25
| Std Dev. |  | 0.21213 | 37
| Rel.Std.Dev. |  | 1.102%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 240b (Indicative Value 2.08 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.08 | 1
| 1 | 2 | 2.07 | 32
| Mean |  | 2.075 | 16
| Median |  | 2.075 | 2.075
| Std Dev. |  | 0.0070711 | 38
| Rel.Std.Dev. |  | 0.3408%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 240b (Indicative Value 173 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 175 | 33
| Mean |  | 172.5 | 16
| Median |  | 172.5 | 172.5
| Std Dev. |  | 3.5355 | 39
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 240b (Indicative Value 70 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 70.5 | 1
| 1 | 2 | 69.5 | 34
| Mean |  | 70 | 16
| Median |  | 70 | 70
| Std Dev. |  | 0.70711 | 40
| Rel.Std.Dev. |  | 1.01%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.