## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 242b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 8.7279 | 8.6495 | 8.8063 | 8.7018 | 8.754
PhotonAssay (recommended gross mass* 358±15 g)
Au, Gold (ppm) | 8.7981 | 8.7211 | 8.875 | 8.7904 | 8.8058
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 8.5188 | 8.3792 | 8.6585 | 8.4909 | 8.5467
Cyanide Leach
Au, Gold (ppm) | 8.4247 | 8.2883 | 8.5612 | 8.415 | 8.4345
4-Acid Digestion
Ag, Silver (ppm) | 4.7499 | 4.5565 | 4.9432 | 4.577 | 4.9227
Al, Aluminium (wt.%) | 6.5315 | 6.3553 | 6.7077 | 6.4171 | 6.6459
As, Arsenic (ppm) | 53.3531 | 50.7852 | 55.921 | 51.6838 | 55.0223
Ba, Barium (ppm) | 152.889 | 146.4624 | 159.3156 | 149.6962 | 156.0818
Be, Beryllium (ppm) | 0.46117 | 0.39686 | 0.52548 | 0.43203 | 0.49031
Bi, Bismuth (ppm) | 0.39182 | 0.3713 | 0.41234 | 0.37594 | 0.40769
Ca, Calcium (wt.%) | 6.3111 | 6.1434 | 6.4787 | 6.2081 | 6.414
Cd, Cadmium (ppm) | 0.39693 | 0.35294 | 0.44092 | 0.36447 | 0.4294
Ce, Cerium (ppm) | 16.9627 | 16.3387 | 17.5868 | 16.5796 | 17.3459
Co, Cobalt (ppm) | 39.6156 | 38.1303 | 41.1008 | 38.7449 | 40.4863
Cr, Chromium (ppm) | 106.9892 | 99.5118 | 114.4667 | 102.5363 | 111.4422
Cs, Caesium (ppm) | 1.0795 | 1.0108 | 1.1481 | 1.0328 | 1.1261
Cu, Copper (ppm) | 173.3828 | 167.8212 | 178.9445 | 169.1026 | 177.663
Dy, Dysprosium (ppm) | 3.8325 | 3.6302 | 4.0348 | 3.7318 | 3.9331
Er, Erbium (ppm) | 2.2975 | 2.1529 | 2.4421 | 2.206 | 2.389
Eu, Europium (ppm) | 0.90098 | 0.81676 | 0.98519 | 0.86018 | 0.94177
Fe, Iron (wt.%) | 7.5002 | 7.2927 | 7.7076 | 7.3957 | 7.6047
Ga, Gallium (ppm) | 15.8051 | 15.1292 | 16.4811 | 15.4607 | 16.1496
Gd, Gadolinium (ppm) | 3.3479 | 3.143 | 3.5529 | 3.2084 | 3.4875
Ge, Germanium (ppm) | 0.12659 | 0.087785 | 0.1654 | IND | IND
Hf, Hafnium (ppm) | 1.7365 | 1.6218 | 1.8513 | 1.6766 | 1.7965
Ho, Holmium (ppm) | 0.76605 | 0.71007 | 0.82202 | 0.73188 | 0.80021
In, Indium (ppm) | 0.07379 | 0.064998 | 0.082582 | 0.06651 | 0.081069
K, Potassium (wt.%) | 0.36313 | 0.34992 | 0.37634 | 0.35532 | 0.37094
La, Lanthanum (ppm) | 7.6023 | 7.2082 | 7.9965 | 7.4293 | 7.7754
Li, Lithium (ppm) | 19.6618 | 18.6751 | 20.6486 | 19.0832 | 20.2404
Lu, Lutetium (ppm) | 0.32405 | 0.29384 | 0.35426 | 0.28771 | 0.3604
Mg, Magnesium (wt.%) | 3.5158 | 3.436 | 3.5956 | 3.4522 | 3.5794
Mn, Manganese (wt.%) | 0.12605 | 0.1213 | 0.13081 | 0.12321 | 0.12889
Mo, Molybdenum (ppm) | 8.1966 | 7.7477 | 8.6454 | 7.8844 | 8.5088
Na, Sodium (wt.%) | 1.7307 | 1.6802 | 1.7811 | 1.689 | 1.7723
Nb, Niobium (ppm) | 3.8943 | 3.7143 | 4.0744 | 3.8018 | 3.9868
Nd, Neodymium (ppm) | 9.9039 | 9.4483 | 10.3595 | 9.7651 | 10.0427
Ni, Nickel (ppm) | 75.2484 | 72.6202 | 77.8766 | 73.6806 | 76.8163
P, Phosphorus (wt.%) | 0.039944 | 0.038221 | 0.041668 | 0.038498 | 0.04139
Pb, Lead (ppm) | 19.0231 | 17.8053 | 20.2409 | 18.4389 | 19.6073
Pr, Praseodymium (ppm) | 2.2059 | 2.0822 | 2.3297 | 2.1155 | 2.2964
Rb, Rubidium (ppm) | 14.6888 | 13.9967 | 15.3809 | 14.3841 | 14.9935
Re, Rhenium (ppm) | 0.0024556 | 0.0014314 | 0.0034797 | IND | IND
S, Sulphur (wt.%) | 0.26461 | 0.25135 | 0.27788 | 0.25734 | 0.27188
Sb, Antimony (ppm) | 3.4569 | 3.2229 | 3.6909 | 3.3147 | 3.5992
Sc, Scandium (ppm) | 37.1853 | 35.3413 | 39.0293 | 36.2152 | 38.1555
Sm, Samarium (ppm) | 2.7846 | 2.5504 | 3.0188 | 2.6758 | 2.8934
Sn, Tin (ppm) | 0.99669 | 0.87727 | 1.1161 | IND | IND
Sr, Strontium (ppm) | 185.2871 | 178.6956 | 191.8787 | 181.0563 | 189.5179
Ta, Tantalum (ppm) | 0.26345 | 0.24351 | 0.2834 | 0.24783 | 0.27908
Tb, Terbium (ppm) | 0.58011 | 0.50331 | 0.65692 | 0.54776 | 0.61247
Te, Tellurium (ppm) | 0.16479 | 0.13292 | 0.19666 | IND | IND
Th, Thorium (ppm) | 1.631 | 1.5399 | 1.722 | 1.5871 | 1.6749
Ti, Titanium (wt.%) | 0.56771 | 0.5515 | 0.58393 | 0.55649 | 0.57894
Tl, Thallium (ppm) | 0.17172 | 0.16029 | 0.18315 | IND | IND
Tm, Thulium (ppm) | 0.31046 | 0.28931 | 0.33161 | 0.28548 | 0.33544
U, Uranium (ppm) | 0.43839 | 0.37903 | 0.49776 | 0.41575 | 0.46104
V, Vanadium (ppm) | 256.8232 | 248.5284 | 265.118 | 249.1853 | 264.4611
W, Tungsten (ppm) | 16.0697 | 15.2467 | 16.8927 | 15.5583 | 16.5811
Y, Yttrium (ppm) | 20.067 | 19.1994 | 20.9347 | 19.5993 | 20.5347
Yb, Ytterbium (ppm) | 2.1629 | 2.0309 | 2.2949 | 2.0909 | 2.235
Zn, Zinc (ppm) | 115.3052 | 110.9524 | 119.6579 | 112.8958 | 117.7145
Zr, Zirconium (ppm) | 54.6677 | 50.4171 | 58.9184 | 51.4699 | 57.8655
Aqua Regia Digestion
Ag, Silver (ppm) | 4.5837 | 4.4365 | 4.7309 | 4.4514 | 4.716
Al, Aluminium (wt.%) | 2.9964 | 2.9017 | 3.091 | 2.939 | 3.0537
As, Arsenic (ppm) | 50.8458 | 48.9551 | 52.7365 | 49.6765 | 52.0151
B, Boron (ppm) | 14.7145 | 11.6619 | 17.7671 | IND | IND
Ba, Barium (ppm) | 30.2786 | 28.6044 | 31.9528 | 29.2086 | 31.3486
Be, Beryllium (ppm) | 0.20375 | 0.17581 | 0.23169 | 0.18867 | 0.21883
Bi, Bismuth (ppm) | 0.39971 | 0.37993 | 0.41949 | 0.38395 | 0.41548
Ca, Calcium (wt.%) | 2.2208 | 1.9819 | 2.4598 | 2.1709 | 2.2708
Cd, Cadmium (ppm) | 0.37325 | 0.34181 | 0.40469 | 0.34459 | 0.40191
Ce, Cerium (ppm) | 10.8423 | 9.904 | 11.7805 | 10.561 | 11.1235
Co, Cobalt (ppm) | 27.3313 | 26.231 | 28.4316 | 26.6707 | 27.992
Cr, Chromium (ppm) | 21.0157 | 19.8697 | 22.1617 | 20.2413 | 21.7902
Cs, Caesium (ppm) | 0.57704 | 0.54068 | 0.61341 | 0.55754 | 0.59654
Cu, Copper (ppm) | 170.2315 | 164.8184 | 175.6446 | 167.0986 | 173.3645
Eu, Europium (ppm) | 0.44474 | 0.32786 | 0.56162 | 0.40982 | 0.47966
Fe, Iron (wt.%) | 5.2873 | 5.1658 | 5.4087 | 5.2113 | 5.3633
Ga, Gallium (ppm) | 9.4956 | 8.959 | 10.0321 | 9.1808 | 9.8103
Ge, Germanium (ppm) | 0.13917 | 0.11042 | 0.16791 | IND | IND
Hf, Hafnium (ppm) | 0.47032 | 0.42304 | 0.5176 | 0.44539 | 0.49525
Hg, Mercury (ppm) | 0.069917 | 0.046926 | 0.092907 | IND | IND
Ho, Holmium (ppm) | 0.43994 | 0.33304 | 0.54683 | 0.39869 | 0.48118
In, Indium (ppm) | 0.032 | 0.029 | 0.035 | 0.029118 | 0.034882
K, Potassium (wt.%) | 0.1085 | 0.10076 | 0.11623 | 0.10391 | 0.11309
La, Lanthanum (ppm) | 3.9156 | 3.6072 | 4.2239 | 3.7704 | 4.0607
Li, Lithium (ppm) | 9.1655 | 8.6177 | 9.7132 | 8.9154 | 9.4156
Mg, Magnesium (wt.%) | 1.5604 | 1.5245 | 1.5964 | 1.5386 | 1.5823
Mn, Manganese (wt.%) | 0.066646 | 0.064903 | 0.068389 | 0.065541 | 0.067751
Mo, Molybdenum (ppm) | 7.8494 | 7.4675 | 8.2314 | 7.6094 | 8.0895
Na, Sodium (wt.%) | 0.1172 | 0.10896 | 0.12544 | 0.11286 | 0.12155
Nb, Niobium (ppm) | 0.11751 | 0.082775 | 0.15224 | IND | IND
Ni, Nickel (ppm) | 48.488 | 46.7488 | 50.2272 | 47.5263 | 49.4496
P, Phosphorus (wt.%) | 0.037803 | 0.036415 | 0.039192 | 0.036867 | 0.03874
Pb, Lead (ppm) | 16.863 | 16.1774 | 17.5485 | 16.4511 | 17.2749
Pr, Praseodymium (ppm) | 1.4079 | 1.0107 | 1.805 | 1.3223 | 1.4934
Rb, Rubidium (ppm) | 6.0256 | 5.6417 | 6.4095 | 5.8041 | 6.2472
Re, Rhenium (ppm) | 0.0015037 | 0.00069972 | 0.0023077 | IND | IND
S, Sulphur (wt.%) | 0.26064 | 0.25026 | 0.27101 | 0.25344 | 0.26783
Sb, Antimony (ppm) | 1.6737 | 1.5027 | 1.8447 | 1.6017 | 1.7458
Sc, Scandium (ppm) | 4.8192 | 4.3259 | 5.3125 | 4.6177 | 5.0208
Se, Selenium (ppm) | 0.87503 | 0.65413 | 1.0959 | IND | IND
Sn, Tin (ppm) | 0.62646 | 0.53518 | 0.71774 | 0.54025 | 0.71266
Sr, Strontium (ppm) | 30.2731 | 27.081 | 33.4651 | 29.3417 | 31.2044
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.30998 | 0.26329 | 0.35668 | 0.28047 | 0.3395
Te, Tellurium (ppm) | 0.14621 | 0.12092 | 0.17149 | IND | IND
Th, Thorium (ppm) | 0.92249 | 0.83151 | 1.0135 | 0.8691 | 0.97587
Ti, Titanium (wt.%) | 0.31904 | 0.28963 | 0.34845 | 0.30791 | 0.33018
Tl, Thallium (ppm) | 0.093992 | 0.084729 | 0.10326 | 0.086718 | 0.10127
U, Uranium (ppm) | 0.24356 | 0.22168 | 0.26544 | 0.22934 | 0.25778
V, Vanadium (ppm) | 124.7271 | 118.1593 | 131.2949 | 121.6523 | 127.8019
W, Tungsten (ppm) | 10.6455 | 9.5806 | 11.7103 | 10.2025 | 11.0884
Y, Yttrium (ppm) | 10.6909 | 10.1569 | 11.225 | 10.3002 | 11.0817
Zn, Zinc (ppm) | 100.5041 | 96.8073 | 104.201 | 97.6548 | 103.3535
Zr, Zirconium (ppm) | 15.9908 | 14.907 | 17.0747 | 15.4604 | 16.5213
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 242b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 39.8333 | Pt | ppb | 13.8333
Hg | ppm | 0.15083 | Se | ppm | 1.1232
Aqua Regia Digestion
Dy | ppm | 2.0148 | Nd | ppm | 5.8304 | Tm | ppm | 0.1679
Er | ppm | 1.15 | Pd | ppb | 15.465 | Yb | ppm | 1.2635
Gd | ppm | 1.8156 | Pt | ppb | 12.3343
Lu | ppm | 0.15259 | Sm | ppm | 1.4414
Borate Fusion XRF
Al2O3 | wt.% | 12.775 | Fe2O3 | wt.% | 10.9587 | S | wt.% | 0.25
As | ppm | 60 | K2O | wt.% | 0.447 | SiO2 | wt.% | 53.795
BaO | ppm | 189.805 | MgO | wt.% | 5.93 | Sn | ppm | 40
CaO | wt.% | 9.185 | MnO | wt.% | 0.171 | Sr | ppm | 200
Cl | ppm | 1250 | Na2O | wt.% | 2.33 | TiO2 | wt.% | 1.016
Co | ppm | 50 | Ni | ppm | 130 | V2O5 | ppm | 482.004
Cr2O3 | ppm | 211.9175 | P2O5 | wt.% | 0.088227 | Zn | ppm | 125
Cu | ppm | 205 | Pb | ppm | 85 | Zr | ppm | 100
Thermogravimetry
LOI1000 | wt.% | 2.965
Infrared Combustion
C | wt.% | 0.14 | S | wt.% | 0.27
Laser Ablation ICP-MS
Ag | ppm | 4.1 | Hf | ppm | 2.12 | Sm | ppm | 2.86
As | ppm | 52.5 | Ho | ppm | 0.835 | Sn | ppm | 0.9
Ba | ppm | 154.5 | In | ppm | 0.075 | Sr | ppm | 183.5
Be | ppm | 0.5 | La | ppm | 7.755 | Ta | ppm | 0.28
Bi | ppm | 0.4 | Lu | ppm | 0.33 | Tb | ppm | 0.58
Cd | ppm | 0.45 | Mn | wt.% | 0.131 | Te | ppm | < 0.2
Ce | ppm | 16.9 | Mo | ppm | 8.6 | Th | ppm | 1.63
Co | ppm | 41.05 | Nb | ppm | 3.975 | Ti | wt.% | 0.597
Cr | ppm | 128.5 | Nd | ppm | 10.1 | Tl | ppm | < 0.2
Cs | ppm | 1.095 | Ni | ppm | 83 | Tm | ppm | 0.35
Cu | ppm | 165 | Pb | ppm | 19 | U | ppm | 0.455
Dy | ppm | 3.88 | Pr | ppm | 2.215 | V | ppm | 276
Er | ppm | 2.335 | Rb | ppm | 14.05 | W | ppm | 16.5
Eu | ppm | 0.83 | Re | ppm | < 0.01 | Y | ppm | 20.7
Ga | ppm | 15.35 | Sb | ppm | 3.5 | Yb | ppm | 2.315
Gd | ppm | 3.355 | Sc | ppm | 36.9 | Zn | ppm | 117.5
Ge | ppm | 1.3 | Se | ppm | < 5 | Zr | ppm | 69.75
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 242b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 8.7279 | 0.19805 | 8.3318 | 9.124 | 8.1337 | 9.3221 | 2.269% | 4.538% | 6.808% | 8.2915 | 9.1643
PhotonAssay
Au, ppm | 8.7981 | 0.1815 | 8.4351 | 9.1611 | 8.2536 | 9.3426 | 2.063% | 4.126% | 6.189% | 8.3582 | 9.238
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 8.5188 | 0.32819 | 7.8624 | 9.1752 | 7.5343 | 9.5034 | 3.852% | 7.705% | 11.56% | 8.0929 | 8.9448
Cyanide Leach
Au, ppm | 8.4247 | 0.31137 | 7.802 | 9.0475 | 7.4906 | 9.3588 | 3.696% | 7.392% | 11.09% | 8.0035 | 8.8459
4-Acid Digestion
Ag, ppm | 4.7499 | 0.15236 | 4.4451 | 5.0546 | 4.2928 | 5.2069 | 3.208% | 6.415% | 9.623% | 4.5124 | 4.9874
Al, wt.% | 6.5315 | 0.25486 | 6.0217 | 7.0412 | 5.7669 | 7.2961 | 3.902% | 7.804% | 11.71% | 6.2049 | 6.858
As, ppm | 53.3531 | 2.8043 | 47.7445 | 58.9616 | 44.9402 | 61.7659 | 5.256% | 10.51% | 15.77% | 50.6854 | 56.0207
Ba, ppm | 152.889 | 5.524 | 141.841 | 163.937 | 136.317 | 169.461 | 3.613% | 7.226% | 10.84% | 145.2446 | 160.5335
Be, ppm | 0.46117 | 0.074294 | 0.31258 | 0.60976 | 0.23829 | 0.68405 | 16.11% | 32.22% | 48.33% | 0.43811 | 0.48423
Bi, ppm | 0.39182 | 0.019797 | 0.35222 | 0.43141 | 0.33243 | 0.45121 | 5.053% | 10.11% | 15.16% | 0.37223 | 0.41141
Ca, wt.% | 6.3111 | 0.24937 | 5.8123 | 6.8098 | 5.563 | 7.0592 | 3.951% | 7.902% | 11.85% | 5.9955 | 6.6266
Cd, ppm | 0.39693 | 0.031739 | 0.33346 | 0.46041 | 0.30172 | 0.49215 | 7.996% | 15.99% | 23.99% | 0.37709 | 0.41678
Ce, ppm | 16.9627 | 0.61788 | 15.727 | 18.1985 | 15.1091 | 18.8164 | 3.643% | 7.285% | 10.93% | 16.1146 | 17.8109
Co, ppm | 39.6156 | 1.391 | 36.8336 | 42.3975 | 35.4426 | 43.7885 | 3.511% | 7.022% | 10.53% | 37.6348 | 41.5964
Cr, ppm | 106.9892 | 8.3805 | 90.2283 | 123.7501 | 81.8479 | 132.1306 | 7.833% | 15.67% | 23.5% | 101.6398 | 112.3387
Cs, ppm | 1.0795 | 0.043416 | 0.99262 | 1.1663 | 0.94921 | 1.2097 | 4.022% | 8.044% | 12.07% | 1.0255 | 1.1334
Cu, ppm | 173.3828 | 5.1483 | 163.0863 | 183.6793 | 157.9381 | 188.8276 | 2.969% | 5.939% | 8.908% | 164.7137 | 182.052
Dy, ppm | 3.8325 | 0.16018 | 3.5121 | 4.1528 | 3.3519 | 4.313 | 4.179% | 8.359% | 12.54% | 3.6408 | 4.0241
Er, ppm | 2.2975 | 0.12466 | 2.0482 | 2.5468 | 1.9235 | 2.6715 | 5.426% | 10.85% | 16.28% | 2.1826 | 2.4124
Eu, ppm | 0.90098 | 0.075233 | 0.75051 | 1.0514 | 0.67528 | 1.1267 | 8.35% | 16.7% | 25.05% | 0.85593 | 0.94602
Fe, wt.% | 7.5002 | 0.28878 | 6.9226 | 8.0777 | 6.6338 | 8.3665 | 3.85% | 7.701% | 11.55% | 7.1252 | 7.8752
Ga, ppm | 15.8051 | 0.70783 | 14.3895 | 17.2208 | 13.6816 | 17.9286 | 4.478% | 8.957% | 13.44% | 15.0149 | 16.5954
Gd, ppm | 3.3479 | 0.13971 | 3.0685 | 3.6274 | 2.9288 | 3.7671 | 4.173% | 8.346% | 12.52% | 3.1806 | 3.5153
Ge, ppm | 0.12659 | 0.04045 | 0.045694 | 0.20749 | 0.0052443 | 0.24794 | 31.95% | 63.9% | 95.86% | 0.12026 | 0.13292
Hf, ppm | 1.7365 | 0.067233 | 1.6021 | 1.871 | 1.5348 | 1.9382 | 3.872% | 7.743% | 11.62% | 1.6497 | 1.8234
Ho, ppm | 0.76605 | 0.034492 | 0.69706 | 0.83503 | 0.66257 | 0.86952 | 4.503% | 9.005% | 13.51% | 0.72774 | 0.80435
In, ppm | 0.07379 | 0.0052657 | 0.063258 | 0.084321 | 0.057993 | 0.089587 | 7.136% | 14.27% | 21.41% | 0.0701 | 0.077479
K, wt.% | 0.36313 | 0.017125 | 0.32888 | 0.39738 | 0.31176 | 0.4145 | 4.716% | 9.432% | 14.15% | 0.34497 | 0.38129
La, ppm | 7.6023 | 0.54795 | 6.5064 | 8.6982 | 5.9585 | 9.2462 | 7.208% | 14.42% | 21.62% | 7.2222 | 7.9825
Li, ppm | 19.6618 | 1.3556 | 16.9507 | 22.373 | 15.5951 | 23.7285 | 6.894% | 13.79% | 20.68% | 18.6787 | 20.6449
Lu, ppm | 0.32405 | 0.013642 | 0.29677 | 0.35133 | 0.28313 | 0.36498 | 4.21% | 8.419% | 12.63% | 0.30785 | 0.34025
Mg, wt.% | 3.5158 | 0.088432 | 3.3389 | 3.6927 | 3.2505 | 3.7811 | 2.515% | 5.031% | 7.546% | 3.34 | 3.6916
Mn, wt.% | 0.12605 | 0.0047174 | 0.11662 | 0.13549 | 0.1119 | 0.14021 | 3.742% | 7.485% | 11.23% | 0.11975 | 0.13236
Mo, ppm | 8.1966 | 0.46715 | 7.2623 | 9.1309 | 6.7951 | 9.598 | 5.699% | 11.4% | 17.1% | 7.7867 | 8.6064
Na, wt.% | 1.7307 | 0.056687 | 1.6173 | 1.844 | 1.5606 | 1.9007 | 3.275% | 6.551% | 9.826% | 1.6441 | 1.8172
Nb, ppm | 3.8943 | 0.1901 | 3.5141 | 4.2745 | 3.324 | 4.4646 | 4.882% | 9.763% | 14.64% | 3.6996 | 4.089
Nd, ppm | 9.9039 | 0.33108 | 9.2417 | 10.566 | 8.9106 | 10.8971 | 3.343% | 6.686% | 10.03% | 9.4087 | 10.3991
Ni, ppm | 75.2484 | 2.8289 | 69.5907 | 80.9062 | 66.7619 | 83.735 | 3.759% | 7.519% | 11.28% | 71.486 | 79.0109
P, wt.% | 0.039944 | 0.0016564 | 0.036632 | 0.043257 | 0.034975 | 0.044914 | 4.147% | 8.294% | 12.44% | 0.037947 | 0.041942
Pb, ppm | 19.0231 | 1.0575 | 16.9081 | 21.1381 | 15.8506 | 22.1956 | 5.559% | 11.12% | 16.68% | 18.0719 | 19.9742
Pr, ppm | 2.2059 | 0.096144 | 2.0137 | 2.3982 | 1.9175 | 2.4944 | 4.358% | 8.717% | 13.08% | 2.0956 | 2.3162
Rb, ppm | 14.6888 | 0.87569 | 12.9374 | 16.4401 | 12.0617 | 17.3158 | 5.962% | 11.92% | 17.88% | 13.9543 | 15.4232
Re, ppm | 0.0024556 | 0.00055238 | 0.0013508 | 0.0035603 | 0.0007984 | 0.0041127 | 22.5% | 44.99% | 67.49% | 0.0023328 | 0.0025783
S, wt.% | 0.26461 | 0.015314 | 0.23398 | 0.29524 | 0.21867 | 0.31055 | 5.787% | 11.57% | 17.36% | 0.25138 | 0.27784
Sb, ppm | 3.4569 | 0.26384 | 2.9292 | 3.9846 | 2.6654 | 4.2485 | 7.632% | 15.26% | 22.9% | 3.2841 | 3.6298
Sc, ppm | 37.1853 | 1.8385 | 33.5084 | 40.8623 | 31.6699 | 42.7007 | 4.944% | 9.888% | 14.83% | 35.3261 | 39.0446
Sm, ppm | 2.7846 | 0.15991 | 2.4648 | 3.1044 | 2.3048 | 3.2643 | 5.743% | 11.49% | 17.23% | 2.6454 | 2.9238
Sn, ppm | 0.99669 | 0.13511 | 0.72647 | 1.2669 | 0.59136 | 1.402 | 13.56% | 27.11% | 40.67% | 0.94686 | 1.0465
Sr, ppm | 185.2871 | 7.944 | 169.3991 | 201.1751 | 161.4551 | 209.1191 | 4.287% | 8.575% | 12.86% | 176.0228 | 194.5515
Ta, ppm | 0.26345 | 0.0091223 | 0.24521 | 0.2817 | 0.23609 | 0.29082 | 3.463% | 6.925% | 10.39% | 0.25028 | 0.27663
Tb, ppm | 0.58011 | 0.063811 | 0.45249 | 0.70773 | 0.38868 | 0.77154 | 11% | 22% | 33% | 0.55111 | 0.60912
Te, ppm | 0.16479 | 0.02832 | 0.10815 | 0.22143 | 0.079833 | 0.24975 | 17.19% | 34.37% | 51.56% | 0.15655 | 0.17303
Th, ppm | 1.631 | 0.11673 | 1.3975 | 1.8644 | 1.2808 | 1.9812 | 7.157% | 14.31% | 21.47% | 1.5494 | 1.7125
Ti, wt.% | 0.56771 | 0.020162 | 0.52739 | 0.60804 | 0.50723 | 0.6282 | 3.551% | 7.103% | 10.65% | 0.53933 | 0.5961
Tl, ppm | 0.17172 | 0.0085825 | 0.15456 | 0.18889 | 0.14597 | 0.19747 | 4.998% | 9.996% | 14.99% | 0.16313 | 0.18031
Tm, ppm | 0.31046 | 0.016075 | 0.27831 | 0.34261 | 0.26223 | 0.35869 | 5.178% | 10.36% | 15.53% | 0.29494 | 0.32598
U, ppm | 0.43839 | 0.039779 | 0.35884 | 0.51795 | 0.31906 | 0.55773 | 9.074% | 18.15% | 27.22% | 0.41647 | 0.46031
V, ppm | 256.8232 | 6.5421 | 243.7391 | 269.9073 | 237.197 | 276.4494 | 2.547% | 5.095% | 7.642% | 243.982 | 269.6643
W, ppm | 16.0697 | 0.74702 | 14.5757 | 17.5637 | 13.8286 | 18.3107 | 4.649% | 9.297% | 13.95% | 15.2662 | 16.8732
Y, ppm | 20.067 | 0.8156 | 18.4358 | 21.6982 | 17.6202 | 22.5138 | 4.064% | 8.129% | 12.19% | 19.0637 | 21.0704
Yb, ppm | 2.1629 | 0.10855 | 1.9458 | 2.38 | 1.8373 | 2.4886 | 5.018% | 10.04% | 15.06% | 2.0548 | 2.2711
Zn, ppm | 115.3052 | 6.9734 | 101.3583 | 129.252 | 94.3849 | 136.2254 | 6.048% | 12.1% | 18.14% | 109.5399 | 121.0704
Zr, ppm | 54.6677 | 3.7513 | 47.1651 | 62.1704 | 43.4137 | 65.9217 | 6.862% | 13.72% | 20.59% | 51.9343 | 57.4011
Aqua Regia Digestion
Ag, ppm | 4.5837 | 0.15618 | 4.2713 | 4.8961 | 4.1152 | 5.0522 | 3.407% | 6.815% | 10.22% | 4.3545 | 4.8129
Al, wt.% | 2.9964 | 0.11633 | 2.7637 | 3.229 | 2.6474 | 3.3453 | 3.882% | 7.764% | 11.65% | 2.8465 | 3.1462
As, ppm | 50.8458 | 2.1945 | 46.4567 | 55.2349 | 44.2622 | 57.4294 | 4.316% | 8.632% | 12.95% | 48.3035 | 53.3881
B, ppm | 14.7145 | 4.5612 | 5.5921 | 23.8369 | 1.0309 | 28.3982 | 31% | 62% | 92.99% | 13.9788 | 15.4502
Ba, ppm | 30.2786 | 2.245 | 25.7886 | 34.7686 | 23.5437 | 37.0135 | 7.414% | 14.83% | 22.24% | 28.7647 | 31.7925
Be, ppm | 0.20375 | 0.026063 | 0.15162 | 0.25588 | 0.12556 | 0.28194 | 12.79% | 25.58% | 38.37% | 0.19356 | 0.21394
Bi, ppm | 0.39971 | 0.018575 | 0.36256 | 0.43686 | 0.34398 | 0.45544 | 4.647% | 9.294% | 13.94% | 0.37972 | 0.4197
Ca, wt.% | 2.2208 | 0.44412 | 1.3326 | 3.1091 | 0.88847 | 3.5532 | 20% | 40% | 59.99% | 2.1098 | 2.3319
Cd, ppm | 0.37325 | 0.028182 | 0.31689 | 0.42961 | 0.2887 | 0.4578 | 7.55% | 15.1% | 22.65% | 0.35459 | 0.39191
Ce, ppm | 10.8423 | 1.3981 | 8.046 | 13.6385 | 6.6479 | 15.0366 | 12.89% | 25.79% | 38.68% | 10.3001 | 11.3844
Co, ppm | 27.3313 | 1.4408 | 24.4498 | 30.2129 | 23.009 | 31.6537 | 5.272% | 10.54% | 15.81% | 25.9648 | 28.6979
Cr, ppm | 21.0157 | 1.6771 | 17.6615 | 24.37 | 15.9843 | 26.0471 | 7.98% | 15.96% | 23.94% | 19.9649 | 22.0665
Cs, ppm | 0.57704 | 0.035598 | 0.50585 | 0.64824 | 0.47025 | 0.68384 | 6.169% | 12.34% | 18.51% | 0.54819 | 0.60589
Cu, ppm | 170.2315 | 5.4412 | 159.3491 | 181.114 | 153.9078 | 186.5552 | 3.196% | 6.393% | 9.589% | 161.7199 | 178.7431
Eu, ppm | 0.44474 | 0.099247 | 0.24625 | 0.64323 | 0.147 | 0.74248 | 22.32% | 44.63% | 66.95% | 0.4225 | 0.46698
Fe, wt.% | 5.2873 | 0.15273 | 4.9818 | 5.5927 | 4.8291 | 5.7455 | 2.889% | 5.777% | 8.666% | 5.0229 | 5.5516
Ga, ppm | 9.4956 | 0.53622 | 8.4231 | 10.568 | 7.8869 | 11.1042 | 5.647% | 11.29% | 16.94% | 9.0208 | 9.9703
Ge, ppm | 0.13917 | 0.021163 | 0.096841 | 0.18149 | 0.075679 | 0.20265 | 15.21% | 30.41% | 45.62% | 0.13221 | 0.14613
Hf, ppm | 0.47032 | 0.040138 | 0.39004 | 0.55059 | 0.3499 | 0.59073 | 8.534% | 17.07% | 25.6% | 0.4468 | 0.49383
Hg, ppm | 0.069917 | 0.014737 | 0.040442 | 0.099391 | 0.025705 | 0.11413 | 21.08% | 42.16% | 63.23% | 0.066421 | 0.073413
Ho, ppm | 0.43994 | 0.067598 | 0.30474 | 0.57513 | 0.23714 | 0.64273 | 15.37% | 30.73% | 46.1% | 0.41794 | 0.46193
In, ppm | 0.032 | 0.0016078 | 0.028784 | 0.035216 | 0.027177 | 0.036823 | 5.024% | 10.05% | 15.07% | 0.0304 | 0.0336
K, wt.% | 0.1085 | 0.0074444 | 0.093608 | 0.12339 | 0.086164 | 0.13083 | 6.861% | 13.72% | 20.58% | 0.10307 | 0.11392
La, ppm | 3.9156 | 0.4546 | 3.0064 | 4.8248 | 2.5518 | 5.2794 | 11.61% | 23.22% | 34.83% | 3.7198 | 4.1113
Li, ppm | 9.1655 | 0.78197 | 7.6015 | 10.7294 | 6.8196 | 11.5114 | 8.532% | 17.06% | 25.6% | 8.7072 | 9.6237
Mg, wt.% | 1.5604 | 0.052608 | 1.4552 | 1.6656 | 1.4026 | 1.7183 | 3.371% | 6.743% | 10.11% | 1.4824 | 1.6385
Mn, wt.% | 0.066646 | 0.0017019 | 0.063242 | 0.07005 | 0.06154 | 0.071752 | 2.554% | 5.107% | 7.661% | 0.063314 | 0.069978
Mo, ppm | 7.8494 | 0.51693 | 6.8156 | 8.8833 | 6.2986 | 9.4002 | 6.586% | 13.17% | 19.76% | 7.457 | 8.2419
Na, wt.% | 0.1172 | 0.009748 | 0.097707 | 0.1367 | 0.087959 | 0.14645 | 8.317% | 16.63% | 24.95% | 0.11134 | 0.12306
Nb, ppm | 0.11751 | 0.030723 | 0.05606 | 0.17895 | 0.025336 | 0.20967 | 26.15% | 52.29% | 78.44% | 0.11163 | 0.12338
Ni, ppm | 48.488 | 2.1992 | 44.0897 | 52.8863 | 41.8905 | 55.0855 | 4.535% | 9.071% | 13.61% | 46.0636 | 50.9124
P, wt.% | 0.037803 | 0.0017044 | 0.034395 | 0.041212 | 0.03269 | 0.042917 | 4.509% | 9.017% | 13.53% | 0.035913 | 0.039694
Pb, ppm | 16.863 | 0.75946 | 15.3441 | 18.3819 | 14.5846 | 19.1414 | 4.504% | 9.007% | 13.51% | 16.0198 | 17.7061
Pr, ppm | 1.4079 | 0.28906 | 0.82974 | 1.986 | 0.54068 | 2.275 | 20.53% | 41.06% | 61.6% | 1.3375 | 1.4782
Rb, ppm | 6.0256 | 0.43957 | 5.1465 | 6.9048 | 4.7069 | 7.3443 | 7.295% | 14.59% | 21.89% | 5.7243 | 6.3269
Re, ppm | 0.0015037 | 0.00049977 | 0.00050417 | 0.0025032 | 4.3977e-06 | 0.003003 | 33.24% | 66.47% | 99.71% | 0.0014285 | 0.0015789
S, wt.% | 0.26064 | 0.011404 | 0.23783 | 0.28344 | 0.22643 | 0.29485 | 4.375% | 8.751% | 13.13% | 0.2476 | 0.27367
Sb, ppm | 1.6737 | 0.26908 | 1.1356 | 2.2119 | 0.86648 | 2.481 | 16.08% | 32.15% | 48.23% | 1.59 | 1.7574
Sc, ppm | 4.8192 | 0.64968 | 3.5199 | 6.1186 | 2.8702 | 6.7683 | 13.48% | 26.96% | 40.44% | 4.5783 | 5.0602
Se, ppm | 0.87503 | 0.17324 | 0.52855 | 1.2215 | 0.35531 | 1.3947 | 19.8% | 39.6% | 59.39% | 0.83128 | 0.91878
Sn, ppm | 0.62646 | 0.094112 | 0.43823 | 0.81468 | 0.34412 | 0.90879 | 15.02% | 30.05% | 45.07% | 0.59513 | 0.65778
Sr, ppm | 30.2731 | 4.8778 | 20.5175 | 40.0286 | 15.6398 | 44.9063 | 16.11% | 32.23% | 48.34% | 28.7594 | 31.7867
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.30998 | 0.034187 | 0.24161 | 0.37836 | 0.20742 | 0.41254 | 11.03% | 22.06% | 33.09% | 0.29448 | 0.32548
Te, ppm | 0.14621 | 0.017823 | 0.11056 | 0.18186 | 0.09274 | 0.19968 | 12.19% | 24.38% | 36.57% | 0.1389 | 0.15352
Th, ppm | 0.92249 | 0.11167 | 0.69915 | 1.1458 | 0.58748 | 1.2575 | 12.11% | 24.21% | 36.32% | 0.87636 | 0.96861
Ti, wt.% | 0.31904 | 0.046926 | 0.22519 | 0.41289 | 0.17827 | 0.45982 | 14.71% | 29.42% | 44.12% | 0.30309 | 0.33499
Tl, ppm | 0.093992 | 0.0073128 | 0.079366 | 0.10862 | 0.072054 | 0.11593 | 7.78% | 15.56% | 23.34% | 0.089293 | 0.098692
U, ppm | 0.24356 | 0.018014 | 0.20754 | 0.27959 | 0.18952 | 0.2976 | 7.396% | 14.79% | 22.19% | 0.23139 | 0.25574
V, ppm | 124.7271 | 8.5938 | 107.5395 | 141.9147 | 98.9457 | 150.5085 | 6.89% | 13.78% | 20.67% | 118.4907 | 130.9635
W, ppm | 10.6455 | 1.54 | 7.5654 | 13.7255 | 6.0253 | 15.2656 | 14.47% | 28.93% | 43.4% | 10.1132 | 11.1777
Y, ppm | 10.6909 | 0.45063 | 9.7897 | 11.5922 | 9.3391 | 12.0428 | 4.215% | 8.43% | 12.65% | 10.1564 | 11.2255
Zn, ppm | 100.5041 | 4.6103 | 91.2835 | 109.7247 | 86.6732 | 114.335 | 4.587% | 9.174% | 13.76% | 95.4789 | 105.5293
Zr, ppm | 15.9908 | 1.2108 | 13.5692 | 18.4125 | 12.3583 | 19.6233 | 7.572% | 15.14% | 22.72% | 15.1913 | 16.7904
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 242b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
NA | Unable to report due to QC failure (Lab 1.15)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 242b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST (SÖGÜT), Sögüt, Bilecik Province, Turkey
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Britannia Mining Solutions, Hamilton, Ontario, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Mineral Solutions (BVMS), Al Wadi District, Jeddah, Saudi Arabia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Limited, Manso Nkwanta, Ashanti Region, Ghana
Intertek Minerals Ltd, Bibiani, Western North Region, Ghana
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Laboratoire LABOMINE SARL, Agadir, Souss-Massa, Morocco
Labwest Minerals Analysis, Perth, WA, Australia
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Ghana Ltd, Obuasi, Ashanti, Ghana
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 242b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 4 | 1 | 267859 | 0.087174 | 8.88 | 8.8186
2 | 1 | ANSLu | 1 | 5 | 3 | 267860 | 0.088513 | 8.8513 | 8.8171
2 | 1 | ANSLu | 1 | 7 | 6 | 267861 | 0.082769 | 8.6649 | 8.8072
2 | 1 | ANSLu | 1 | 3 | 10 | 267862 | 0.082997 | 8.7547 | 8.812
2 | 1 | ANSLu | 1 | 17 | 12 | 267863 | 0.082481 | 8.8271 | 8.8158
2 | 1 | ANSLu | 1 | 18 | 7 | 267864 | 0.08366 | 8.7124 | 8.8097
2 | 1 | ANSLu | 1 | 13 | 3 | 267865 | 0.08277 | 8.5074 | 8.7988
2 | 1 | ANSLu | 1 | 6 | 4 | 267866 | 0.084081 | 8.6986 | 8.809
2 | 1 | ANSLu | 1 | 19 | 10 | 267867 | 0.08569 | 8.9129 | 8.8204
2 | 1 | ANSLu | 1 | 9 | 9 | 267868 | 0.086817 | 8.8711 | 8.8182
2 | 1 | ANSLu | 1 | 20 | 11 | 267869 | 0.084997 | 8.9819 | 8.824
2 | 1 | ANSLu | 1 | 8 | 2 | 267870 | 0.088226 | 8.789 | 8.8138
2 | 1 | ANSLu | 1 | 2 | 9 | 267871 | 0.08358 | 8.9294 | 8.8213
2 | 1 | ANSLu | 1 | 1 | 4 | 267872 | 0.083942 | 8.4849 | 8.7976
2 | 1 | ANSLu | 1 | 11 | 1 | 267873 | 0.082929 | 9.0485 | 8.8276
2 | 1 | ANSLu | 1 | 10 | 12 | 267874 | 0.084919 | 8.7921 | 8.814
2 | 1 | ANSLu | 1 | 14 | 5 | 267875 | 0.087532 | 8.9116 | 8.8203
2 | 1 | ANSLu | 1 | 12 | 8 | 267876 | 0.083276 | 9.0535 | 8.8279
2 | 1 | ANSLu | 1 | 15 | 7 | 267877 | 0.0842 | 8.9384 | 8.8217
2 | 1 | ANSLu | 1 | 16 | 6 | 267878 | 0.085842 | 8.6937 | 8.8087
Mean | [6 empty] | 0.08482 | 8.8152 | 8.8152
Median | [7 empty] | 8.8392 | 8.8165
SD | [7 empty] | 0.15645 | 0.0083189
RSD | [7 empty] | 1.775% | 0.09437%
RSD Upscale Factor | [7 empty] | 0.053173
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 242b (Certified Value 8.73 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.28 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 50g | 25g | 30g | 30g | 50g | 50g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 40g | 50g | 30g | 40g | 3
| 1 | 1 | 8.4849 | 8.91 | 8.936 | 8.5562 | 8.934 | 8.4235 | 8.9875 | 8.43 | 8.77 | 8.72 | 8.59 | 8.67 | 8.55 | 8.67 | 8.8729 | 8.66 | 9.16 | 8.656 | 8.306 | 8.635 | 8.97 | 8.83 | 8.646 | 8.7 | 1
| 1 | 2 | 8.9294 | 8.87 | 8.846 | 8.5526 | 8.58 | 8.5312 | 8.7225 | 8.65 | 8.83 | 8.82 | 8.71 | 8.69 | 8.4 | 8.52 | 9.1142 | 8.67 | 9.15 | 8.558 | 8.603 | 8.87 | 9.61 | 8.892 | 8.672 | 8.7 | #N/A
| 1 | 3 | 8.7547 | 8.71 | 8.752 | 8.5852 | 8.632 | 8.601 | 8.6875 | 8.63 | 8.85 | 8.97 | 8.71 | 8.66 | 8.55 | 8.62 | 8.7001 | 8.86 | 9.1 | 8.436 | 8.502 | 8.985 | 9.25 | 8.951 | 8.576 | 8.7 | 16
| 1 | 4 | 8.88 | 8.93 | 8.675 | 8.5643 | 8.913 | 8.5198 | 8.84 | 8.49 | 8.87 | 8.82 | 8.7 | 8.65 | 8.63 | 8.58 | 8.9894 | 8.72 | 9.09 | 8.385 | 8.323 | 8.81 | 8.91 | 9.005 | 8.622 | 8.6 | 8.7279
| 1 | 5 | 8.8513 | 9 | 8.948 | 8.5936 | 8.722 | 8.4506 | 8.715 | 8.69 | 8.82 | 8.88 | 8.61 | 8.87 | 8.49 | 8.68 | 9.1446 | 8.66 | 9.09 | 8.504 | 8.512 | 8.558 | 8.89 | 8.975 | 8.525 | 8.6 | 7
| 1 | 6 | 8.6986 | 8.81 | 8.871 | 8.5518 | 8.852 | 8.5971 | 8.62 | 8.3 | 8.78 | 8.8 | 8.6 | 8.6 | 8.55 | 8.52 | 8.7125 | 8.74 | 9.11 | 8.494 | 8.431 | 8.903 | 8.9 | 9.162 | 8.645 | 8.8
|  | 7 | 8.6649
|  | 8 | 8.789
|  | 9 | 8.8711
|  | 10 | 8.7921
|  | 11 | 9.0485
|  | 12 | 9.0535
|  | 13 | 8.5074
|  | 14 | 8.9116
|  | 15 | 8.9384
|  | 16 | 8.6937
|  | 17 | 8.8271
|  | 18 | 8.7124
|  | 19 | 8.9129
|  | 20 | 8.9819
| Mean |  | 8.8152 | 8.8717 | 8.838 | 8.5673 | 8.7722 | 8.5205 | 8.7621 | 8.5317 | 8.82 | 8.835 | 8.6533 | 8.69 | 8.5283 | 8.5983 | 8.9223 | 8.7183 | 9.1167 | 8.5055 | 8.4462 | 8.7935 | 9.0883 | 8.9692 | 8.6143 | 8.6833
| Median |  | 8.8392 | 8.89 | 8.8585 | 8.5603 | 8.787 | 8.5255 | 8.7188 | 8.56 | 8.825 | 8.82 | 8.655 | 8.665 | 8.55 | 8.6 | 8.9312 | 8.695 | 9.105 | 8.499 | 8.4665 | 8.84 | 8.94 | 8.963 | 8.6335 | 8.7
| Std Dev. |  | 0.15645 | 0.10128 | 0.1066 | 0.017893 | 0.14932 | 0.07315 | 0.13147 | 0.15132 | 0.038987 | 0.083845 | 0.058878 | 0.093167 | 0.07705 | 0.070545 | 0.19317 | 0.07705 | 0.030768 | 0.094754 | 0.11584 | 0.16454 | 0.28944 | 0.11327 | 0.054364 | 0.075277
| Rel.Std.Dev. |  | 1.775% | 1.142% | 1.206% | 0.2089% | 1.702% | 0.8585% | 1.5% | 1.774% | 0.442% | 0.949% | 0.6804% | 1.072% | 0.9035% | 0.8205% | 2.165% | 0.8838% | 0.3375% | 1.114% | 1.372% | 1.871% | 3.185% | 1.263% | 0.6311% | 0.8669%
| PDM3 |  | 1% | 1.647% | 1.261% | -1.84% | 0.5072% | -2.376% | 0.3917% | -2.248% | 1.055% | 1.227% | -0.8543% | -0.4342% | -2.287% | -1.485% | 2.227% | -0.1096% | 4.454% | -2.548% | -3.228% | 0.7516% | 4.13% | 2.764% | -1.301% | -0.5106%
| Z-Score (Absolute) |  | NA | 0.86 | 0.67 | 0.85 | 0.3 | 1.11 | 0.25 | 1.05 | 0.57 | 0.66 | 0.37 | 0.16 | 1.07 | 0.67 | 1.15 | 0 | 2.24 | 1.2 | 1.53 | 0.42 | 2.08 | 1.41 | 0.58 | 0.2
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 242b (Certified Value 8.8 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | 6.13 | 6.14 | 6.15 | 6.16 | 6.17 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 358±15 g | 3
| 1 | 1 | 8.4849 | 8.66 | 9.02 | 8.644 | 8.48 | 8.65 | 8.891 | 8.879 | 8.767 | 8.897 | 8.78 | 8.984 | 8.592 | 8.582 | 8.746 | 8.935 | 8.85 | 8.89 | 1
| 1 | 2 | 8.9294 | 9.09 | 8.98 | 8.461 | 8.55 | 8.71 | 8.762 | 9.019 | 8.783 | 8.758 | 8.93 | 8.92 | 8.63 | 8.545 | 8.98 | 8.935 | 9.04 | 8.992 | #N/A
| 1 | 3 | 8.7547 | [16 empty] | 8.748 | 16
| 1 | 4 | 8.88 | [16 empty] | 9.049 | 8.7981
| 1 | 5 | 8.8513 | 8.65 | 8.84 | 8.505 | 8.37 | 9.02 | 8.832 | 9.135 | 8.907 | 8.799 | 8.88 | 8.769 | 8.632 | 8.455 | 8.85 | 9.111 | 8.99 | 8.862 | 9
| 1 | 6 | 8.6986 | 8.74 | 9 | 8.821 | 8.42 | 8.56 | 8.858 | 9.035 | 8.748 | 8.693 | 8.75 | 8.778 | 8.563 | 8.656 | 8.702 | 9.111 | 8.97 | 8.915
| 1 | 7 | 8.6649 | [16 empty] | 8.758
| 1 | 8 | 8.789 | [16 empty] | 8.875
| 1 | 9 | 8.8711 | 8.6 | 8.93 | 8.646 | 8.66 | 8.87 | 9.191 | 9.136 | 8.839 | 9.029 | 8.97 | 9.03 | 8.542 | 8.54 | 8.638 | 8.889 | 8.74 | 8.697
| 1 | 10 | 8.7921 | 8.76 | 8.85 | 8.525 | 8.58 | 8.82 | 8.667 | 8.914 | 8.961 | 8.808 | 8.93 | 8.771 | 8.547 | 8.704 | 8.717 | 8.889 | 8.87 | 8.791
| 1 | 11 | 9.0485 | [16 empty] | 8.741
| 1 | 12 | 9.0535 | [16 empty] | 9.163
|  | 13 | 8.5074
|  | 14 | 8.9116
|  | 15 | 8.9384
|  | 16 | 8.6937
|  | 17 | 8.8271
|  | 18 | 8.7124
|  | 19 | 8.9129
|  | 20 | 8.9819
| Mean |  | 8.8152 | 8.75 | 8.9367 | 8.6003 | 8.51 | 8.7717 | 8.8668 | 9.0197 | 8.8342 | 8.8307 | 8.8733 | 8.8753 | 8.5843 | 8.5803 | 8.7722 | 8.9783 | 8.91 | 8.8734
| Median |  | 8.8392 | 8.7 | 8.955 | 8.5845 | 8.515 | 8.765 | 8.845 | 9.027 | 8.811 | 8.8035 | 8.905 | 8.849 | 8.5775 | 8.5635 | 8.7315 | 8.935 | 8.92 | 8.8685
| Std Dev. |  | 0.15645 | 0.17686 | 0.077115 | 0.13181 | 0.10733 | 0.16558 | 0.1777 | 0.10772 | 0.084816 | 0.11789 | 0.089144 | 0.11781 | 0.040143 | 0.089 | 0.12321 | 0.1048 | 0.11045 | 0.13993
| Rel.Std.Dev. |  | 1.775% | 2.021% | 0.8629% | 1.533% | 1.261% | 1.888% | 2.004% | 1.194% | 0.9601% | 1.335% | 1.005% | 1.327% | 0.4676% | 1.037% | 1.405% | 1.167% | 1.24% | 1.577%
| PDM3 |  | 0.1944% | -0.5464% | 1.575% | -2.248% | -3.274% | -0.3001% | 0.7815% | 2.519% | 0.4102% | 0.3705% | 0.8554% | 0.8781% | -2.429% | -2.475% | -0.2945% | 2.049% | 1.272% | 0.8564%
| Z-Score (Absolute) |  | NA | 0.75 | 0.91 | 2.08 | 2.88 | 0.56 | 0.29 | 1.65 | 0 | 0.03 | 0.35 | 0.37 | 2.22 | 2.26 | 0.55 | 1.28 | 0.67 | 0.35
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 242b (Certified Value 8.52 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*OES | AR*AAS | AR*AAS | AR*MS | AR*AAS | 2
|  |  | 0.085g | 50g | 25g | 30g | 50g | 30g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 10g | 30g | 15g | 10g | 15g | 25g | 30g | 10g | 40g | 25g | 3
| 1 | 1 | 8.4849 | 8.69 | 8.39 | 8.5273 | 8.9 | 8.1 | 8.61 | 9.4 | 8.72 | 8.38 | 8.56 | 8.7 | 8.8486 | 7.935 | 8.53 | 8.2214 | > 2 | 8.2325 | 7.71 | 8.54 | 8.87 | 8.12 | 8.29 | 1
| 1 | 2 | 8.9294 | 8.57 | 8.65 | 8.4925 | 8.89 | 7.44 | 8.75 | 8.91 | 8.81 | 8.43 | 8.52 | 8.71 | 8.9037 | 8.112 | 8.6 | 8.1709 | > 2 | 8.0951 | 9.11 | 8.29 | 8.21 | 8.2 | 8.27 | #N/A
| 1 | 3 | 8.7547 | 8.94 | 8.58 | 8.5372 | 8.4075 | 7.74 | 8.76 | 9.07 | 8.78 | 8.64 | 8.5 | 8.86 | 8.9228 | 8.389 | 8.52 | 8.3123 | > 2 | 7.9656 | 9.14 | 8.26 | 8.69 | 8.24 | 8.28 | 16
| 1 | 4 | 8.88 | 8.81 | 8.69 | 8.5726 | 8.38 | 7.78 | 8.74 | 8.65 | 8.82 | 9.02 | 8.6 | 8.72 | 8.9754 | 7.759 | 8.7 | 8.1204 | > 2 | 7.79 | 8.8 | 8.44 | 8.88 | 8.3 | 8.29 | 8.5188
| 1 | 5 | 8.8513 | 8.69 | 8.51 | 8.472 | 8.3875 | 8.19 | 8.79 | 8.68 | 8.79 | 8.42 | 8.57 | 8.86 | 8.9459 | 8.219 | 8.52 | 8.2845 | > 2 | 7.9359 | 8.7 | 8.31 | 8.89 | 8.33 | 8.26 | 11
| 1 | 6 | 8.6986 | 8.76 | 8.64 | 8.5718 | 8.4325 | 7.76 | 8.79 | 8.6 | 8.62 | 8.54 | 8.67 | 8.78 | 8.8265 | 8.019 | 8.61 | 8.337 | > 2 | 7.8161 | 8.84 | 8.07 | 8.96 | 8.27 | 8.24
|  | 7 | 8.6649
|  | 8 | 8.789
|  | 9 | 8.8711
|  | 10 | 8.7921
|  | 11 | 9.0485
|  | 12 | 9.0535
|  | 13 | 8.5074
|  | 14 | 8.9116
|  | 15 | 8.9384
|  | 16 | 8.6937
|  | 17 | 8.8271
|  | 18 | 8.7124
|  | 19 | 8.9129
|  | 20 | 8.9819
| Mean |  | 8.8152 | 8.7433 | 8.5767 | 8.5289 | 8.5663 | 7.835 | 8.74 | 8.885 | 8.7567 | 8.5717 | 8.57 | 8.7717 | 8.9038 | 8.0722 | 8.58 | 8.2411 |  | 7.9725 | 8.7167 | 8.3183 | 8.75 | 8.2433 | 8.2717
| Median |  | 8.8392 | 8.725 | 8.61 | 8.5322 | 8.42 | 7.77 | 8.755 | 8.795 | 8.785 | 8.485 | 8.565 | 8.75 | 8.9133 | 8.0655 | 8.565 | 8.2529 |  | 7.9507 | 8.82 | 8.3 | 8.875 | 8.255 | 8.275
| Std Dev. |  | 0.15645 | 0.1258 | 0.11094 | 0.040958 | 0.25532 | 0.27201 | 0.066933 | 0.30937 | 0.075542 | 0.23937 | 0.060663 | 0.073869 | 0.057057 | 0.22057 | 0.071274 | 0.084893 |  | 0.1684 | 0.5234 | 0.16117 | 0.27936 | 0.075542 | 0.019408
| Rel.Std.Dev. |  | 1.775% | 1.439% | 1.293% | 0.4802% | 2.981% | 3.472% | 0.7658% | 3.482% | 0.8627% | 2.793% | 0.7079% | 0.8421% | 0.6408% | 2.732% | 0.8307% | 1.03% |  | 2.112% | 6.005% | 1.938% | 3.193% | 0.9164% | 0.2346%
| PDM3 |  | 3.479% | 2.636% | 0.6791% | 0.1183% | 0.5568% | -8.027% | 2.596% | 4.299% | 2.792% | 0.6204% | 0.6009% | 2.968% | 4.52% | -5.243% | 0.7183% | -3.26% |  | -6.413% | 2.323% | -2.353% | 2.714% | -3.234% | -2.901%
| Z-Score (Absolute) |  | NA | 0.63 | 0.02 | 0.16 | 0.02 | 2.69 | 0.61 | 1.14 | 0.67 | 0 | 0.01 | 0.73 | 1.21 | 1.82 | 0.03 | 1.2 | NA | 2.18 | 0.53 | 0.92 | 0.65 | 1.2 | 1.09
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 242b (Certified Value 8.42 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.22 | 1.23 | 1.27 | 1.29 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 20g | 15g | 05g | 30g | 200g | 50g | 200g | 50g | 3
| 1 | 1 | 8.4849 | 8.94 | 8.25 | 7.948 | 8.563 | 8.3983 | 8.14 | 8.59 | 8.7 | 8.69 | 8.24 | 8.28 | 8.67 | 8.2169 | 8.06 | 7.8 | 8.13 | 8.479 | 8.595 | 8.2 | 8.61 | 1
| 1 | 2 | 8.9294 | 8.77 | 8.44 | 8.0528 | 8.623 | 8.5395 | 8.64 | 8.51 | 8.69 | 8.71 | 8.37 | 8.07 | 8.84 | 8.1214 | 8.11 | 7.5 | 8.4 | 8.715 | 8.726 | 8.65 | 8.52 | #N/A
| 1 | 3 | 8.7547 | 8.85 | 8.39 | 7.9175 | 8.706 | 8.4757 | 7.95 | 8.97 | 8.73 | 8.79 | 8.44 | 8.26 | 8.76 | 7.9538 | 8.4 | 7.6 | 8.4 | 8.533 | 8.575 | 8.85 | 8.5 | 16
| 1 | 4 | 8.88 | 8.75 | 8.2 | 8.0293 | 8.593 | 8.4277 | 8.04 | 8.93 | 8.71 | 8.88 | 8.34 | 8.1 | 8.7 | 8.1214 | 8.08 | 7.9 | 8.39 | 8.792 | 8.565 | 8.76 | 8.56 | 8.4247
| 1 | 5 | 8.8513 | 8.79 | 8.1 | 7.9142 | 8.607 | 8.3476 | 8.31 | 8.86 | 8.71 | 8.9 | 8.37 | 8.34 | 8.64 | 8.2131 | 8.23 | 7.8 | 8.37 | 8.221 | 8.599 | 8.5 | 8.79 | 13
| 1 | 6 | 8.6986 | 8.68 | 8.23 | 7.8639 | 8.596 | 8.4182 | 8.09 | 8.73 | 8.68 | 8.63 | 8.49 | 8.33 | 8.74 | 7.9992 | 9.45 | 8 | 8.34 | 8.237 | 8.546 | 8.57 | 8.59
|  | 7 | 8.6649
|  | 8 | 8.789
|  | 9 | 8.8711
|  | 10 | 8.7921
|  | 11 | 9.0485
|  | 12 | 9.0535
|  | 13 | 8.5074
|  | 14 | 8.9116
|  | 15 | 8.9384
|  | 16 | 8.6937
|  | 17 | 8.8271
|  | 18 | 8.7124
|  | 19 | 8.9129
|  | 20 | 8.9819
| Mean |  | 8.8152 | 8.7967 | 8.2683 | 7.9543 | 8.6147 | 8.4345 | 8.195 | 8.765 | 8.7033 | 8.7667 | 8.375 | 8.23 | 8.725 | 8.1043 | 8.3883 | 7.7667 | 8.3383 | 8.4962 | 8.601 | 8.5883 | 8.595
| Median |  | 8.8392 | 8.78 | 8.24 | 7.9327 | 8.6015 | 8.423 | 8.115 | 8.795 | 8.705 | 8.75 | 8.37 | 8.27 | 8.72 | 8.1214 | 8.17 | 7.8 | 8.38 | 8.506 | 8.585 | 8.61 | 8.575
| Std Dev. |  | 0.15645 | 0.089368 | 0.12576 | 0.072788 | 0.048903 | 0.06618 | 0.24874 | 0.18716 | 0.017512 | 0.10857 | 0.085965 | 0.11662 | 0.071484 | 0.10842 | 0.53529 | 0.18619 | 0.10458 | 0.23662 | 0.064284 | 0.22833 | 0.10407
| Rel.Std.Dev. |  | 1.775% | 1.016% | 1.521% | 0.9151% | 0.5677% | 0.7846% | 3.035% | 2.135% | 0.2012% | 1.238% | 1.026% | 1.417% | 0.8193% | 1.338% | 6.381% | 2.397% | 1.254% | 2.785% | 0.7474% | 2.659% | 1.211%
| PDM3 |  | 4.635% | 4.415% | -1.856% | -5.584% | 2.255% | 0.1164% | -2.727% | 4.039% | 3.307% | 4.059% | -0.5901% | -2.311% | 3.564% | -3.804% | -0.4318% | -7.811% | -1.025% | 0.8481% | 2.093% | 1.942% | 2.021%
| Z-Score (Absolute) |  | NA | 1.03 | 0.61 | 1.59 | 0.47 | 0.1 | 0.84 | 0.93 | 0.74 | 0.94 | 0.28 | 0.73 | 0.81 | 1.13 | 0.24 | 2.18 | 0.4 | 0.1 | 0.42 | 0.38 | 0.4
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 242b (Certified Value 4.75 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 4.8 | 4.6957 | 4.6 | 5 | 4.76 | 4.8 | 4.83 | 4.42 | 4.77 | 5.0345 | 4.4 | 5.1 | 4.802 | 4.98 | 4.483 | 4.8 | 6.65 | 5 | 5.079 |  | 1
| 1 | 2 | 4.84 | 4.7223 | 4.6 | 4.7 | 4.85 | 4.72 | 4.68 | 4.68 | 4.7 | 5.2482 | 4.4 | 5.2 | 4.657 | 5.01 | 4.473 | 4.5 | 6.65 | 4.8 | 6.864 |  | 22
| 1 | 3 | 4.99 | 4.9489 | 4.6 | 4.7 | 4.7 | 4.79 | 4.8 | 4.89 | 4.73 | 5.0148 | 4.5 | 5.2 | 4.67 | 5.13 | 4.535 | 4.8 | 6.31 | 4.8 | 0.763 |  | 16
| 1 | 4 | 5.01 | 4.5889 | 4.7 | 4.8 | 4.58 | 4.69 | 4.72 | 4.81 | 4.72 | 5.2417 | 4.4 | 5.2 | 4.883 | 4.95 | 4.542 | 4.8 | 6.22 | 4.8 | 5.015 |  | 4.7499
| 1 | 5 | 4.85 | 4.5037 | 4.8 | 4.7 | 4.68 | 4.75 | 4.44 | 4.75 | 4.71 | 5.1746 | 4.4 | 5.1 | 4.76 | 5.03 | 4.505 | 4.7 | 6.25 | 5 | 7.029 |  | 15
| 1 | 6 | 4.94 | 4.7546 | 4.7 | 4.7 | 4.71 | 4.74 | 4.72 | 4.48 | 4.54 | 5.099 | 4.4 | 5.1 | 4.445 | 4.96 | 4.539 | 4.9 | 6.66 | 5 | 4.863
| Mean |  | 4.905 | 4.7023 | 4.6667 | 4.7667 | 4.7133 | 4.7483 | 4.6983 | 4.6717 | 4.695 | 5.1355 | 4.4167 | 5.15 | 4.7028 | 5.01 | 4.5128 | 4.75 | 6.4567 | 4.9 | 4.9355
| Median |  | 4.895 | 4.709 | 4.65 | 4.7 | 4.705 | 4.745 | 4.72 | 4.715 | 4.715 | 5.1368 | 4.4 | 5.15 | 4.715 | 4.995 | 4.52 | 4.8 | 6.48 | 4.9 | 5.047
| Std Dev. |  | 0.086891 | 0.15253 | 0.08165 | 0.12111 | 0.089368 | 0.041673 | 0.13834 | 0.18606 | 0.079687 | 0.1016 | 0.040825 | 0.054772 | 0.1518 | 0.06603 | 0.030215 | 0.13784 | 0.21741 | 0.10954 | 2.2603
| Rel.Std.Dev. |  | 1.771% | 3.244% | 1.75% | 2.541% | 1.896% | 0.8776% | 2.944% | 3.983% | 1.697% | 1.978% | 0.9243% | 1.064% | 3.228% | 1.318% | 0.6695% | 2.902% | 3.367% | 2.236% | 45.8%
| PDM3 |  | 3.266% | -1% | -1.751% | 0.3539% | -0.769% | -0.0321% | -1.085% | -1.646% | -1.155% | 8.119% | -7.015% | 8.424% | -0.99% | 5.477% | -4.99% | 0.002993% | 35.93% | 3.161% | 3.908%
| Z-Score (Absolute) |  | 1.29 | 0.38 | 0.67 | 0.15 | 0.29 | 0 | 0.41 | 0.63 | 0.44 | 3.2 | 2.74 | 3.32 | 0.38 | 2.16 | 1.94 | 0.01 | 14.11 | 1.25 | 1.55
| Analytical results for Al in OREAS 242b (Certified Value 6.53 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 6.8475 | 6.6246 | 6.8671 | 6.33 | 6.66 | 6.58 | 6.28 | 6.48 | 6.3 | 6.55 | 6.8452 | 6.2429 | 6.77 | 6.74 | 5.9701 | 6.27 | 7.05 | 6.4 | 6.35 | 6.7313 | 1
| 1 | 2 | 6.7761 | 6.4749 | 6.885 | 6.22 | 6.66 | 6.51 | 6.45 | 6.4 | 6.56 | 6.6 | 6.9405 | 6.3905 | 6.97 | 6.82 | 5.9619 | 6.61 | 6.61 | 6.21 | 6.31 | 6.7251 | #N/A
| 1 | 3 | 6.7949 | 6.7717 | 6.8823 | 6.25 | 6.51 | 6.5 | 6.38 | 6.3 | 6.6 | 6.59 | 6.8681 | 6.3622 | 6.8 | 6.43 | 6.1103 | 6.36 | 6.23 | 6.34 | 6.3 | 7.027 | 16
| 1 | 4 | 6.8659 | 6.534 | 6.8937 | 6.34 | 6.64 | 6.49 | 6.24 | 6.26 | 6.65 | 6.55 | 6.8368 | 6.2803 | 6.86 | 6.94 | 6.0007 | 6.42 | 6.69 | 6.36 | 6.3 | 6.7796 | 6.5315
| 1 | 5 | 6.9713 | 6.5464 | 6.8715 | 6.34 | 6.66 | 6.5 | 6.52 | 6.22 | 6.59 | 6.45 | 6.8008 | 6.2366 | 6.86 | 6.71 | 6.0319 | 6.48 | 6.02 | 6.3 | 6.35 | 6.6505 | 16
| 1 | 6 | 6.8755 | 6.6916 | 6.8891 | 6.28 | 6.54 | 6.59 | 6.4 | 6.36 | 6.64 | 6.46 | 6.9292 | 6.3997 | 6.79 | 6.48 | 5.9327 | 6.4 | 6.45 | 6.3 | 6.33 | 6.6339
| Mean |  | 6.8552 | 6.6072 | 6.8815 | 6.2933 | 6.6117 | 6.5283 | 6.3783 | 6.3367 | 6.5567 | 6.5333 | 6.8701 | 6.3187 | 6.8417 | 6.6867 | 6.0013 | 6.4233 | 6.5083 | 6.3183 | 6.3233 | 6.7579
| Median |  | 6.8567 | 6.5855 | 6.8837 | 6.305 | 6.65 | 6.505 | 6.39 | 6.33 | 6.595 | 6.55 | 6.8567 | 6.3212 | 6.83 | 6.725 | 5.9854 | 6.41 | 6.53 | 6.32 | 6.32 | 6.7282
| Std Dev. |  | 0.069221 | 0.11051 | 0.010269 | 0.051251 | 0.06824 | 0.04446 | 0.10439 | 0.095847 | 0.13003 | 0.064083 | 0.054762 | 0.074255 | 0.073052 | 0.19694 | 0.063324 | 0.11501 | 0.36224 | 0.065243 | 0.023381 | 0.14251
| Rel.Std.Dev. |  | 1.01% | 1.673% | 0.1492% | 0.8144% | 1.032% | 0.681% | 1.637% | 1.513% | 1.983% | 0.9809% | 0.7971% | 1.175% | 1.068% | 2.945% | 1.055% | 1.79% | 5.566% | 1.033% | 0.3698% | 2.109%
| PDM3 |  | 4.957% | 1.159% | 5.358% | -3.646% | 1.228% | -0.04797% | -2.345% | -2.982% | 0.3858% | 0.02858% | 5.185% | -3.258% | 4.749% | 2.376% | -8.118% | -1.656% | -0.3542% | -3.263% | -3.187% | 3.467%
| Z-Score (Absolute) |  | 1.09 | 0.26 | 1.18 | 0.8 | 0.27 | 0.01 | 0.51 | 0.65 | 0.09 | 0.01 | 1.14 | 0.71 | 1.04 | 0.52 | 1.78 | 0.36 | 0.08 | 0.71 | 0.7 | 0.76
| Analytical results for As in OREAS 242b (Certified Value 53 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 0
1
| 1 | 1 | 53.5 | 53.9839 | 55.8 | 57 | 57.9 | 52.8 | 56.6 | 54.1 | 54.6 | 50.9134 | 42.1 | 53 | 51.7 | 50.9 | 44.4 | 52.9 | 32.41 | 54 | 104.661 |  | 1
| 1 | 2 | 54 | 53.3081 | 56.5 | 54 | 58.3 | 52.8 | 55.6 | 60.8 | 56.9 | 51.5349 | 43.8 | 55 | 52.2 | 50.1 | 47.1 | 51.5 | 31.14 | 53 | 95.548 |  | 10
| 1 | 3 | 55.7 | 54.9449 | 55.8 | 54 | 56.1 | 54 | 56.8 | 57 | 52.9 | 53.0855 | 43.6 | 54 | 49.2 | 52.2 | 45.2 | 54 | 32.36 | 54 | 226.894 |  | 16
| 1 | 4 | 54.7 | 53.8976 | 56.2 | 53 | 55.3 | 52.3 | 55 | 57.4 | 54.8 | 52.8579 | 42.8 | 53 | 52.3 | 50.3 | 46.4 | 52.7 | 32.08 | 54 | 100.884 |  | 53.3531
| 1 | 5 | 55.7 | 53.0402 | 56.1 | 54 | 57 | 50.6 | 53.3 | 56.1 | 52.6 | 49.6686 | 42.3 | 54 | 49.8 | 51.3 | 46.3 | 53 | 31.57 | 54 | 94.81 |  | 17
| 1 | 6 | 54.8 | 54.1716 | 56.1 | 55 | 56.9 | 54 | 55.8 | 54.5 | 53.1 | 49.5077 | 44.6 | 54 | 47.8 | 50.3 | 45.4 | 52.5 | 32.53 | 53 | 96.169
| Mean |  | 54.7333 | 53.891 | 56.0833 | 54.5 | 56.9167 | 52.75 | 55.5167 | 56.65 | 54.15 | 51.2613 | 43.2 | 53.8333 | 50.5 | 50.85 | 45.8 | 52.7667 | 32.015 | 53.6667 | 119.8277
| Median |  | 54.75 | 53.9407 | 56.1 | 54 | 56.95 | 52.8 | 55.7 | 56.55 | 53.85 | 51.2241 | 43.2 | 54 | 50.75 | 50.6 | 45.85 | 52.8 | 32.22 | 54 | 98.5265
| Std Dev. |  | 0.88694 | 0.67258 | 0.26394 | 1.3784 | 1.1107 | 1.2613 | 1.2719 | 2.4222 | 1.6282 | 1.5287 | 0.9654 | 0.75277 | 1.8461 | 0.79937 | 0.97775 | 0.80911 | 0.54921 | 0.5164 | 52.5875
| Rel.Std.Dev. |  | 1.62% | 1.248% | 0.4706% | 2.529% | 1.951% | 2.391% | 2.291% | 4.276% | 3.007% | 2.982% | 2.235% | 1.398% | 3.656% | 1.572% | 2.135% | 1.533% | 1.715% | 0.9622% | 43.89%
| PDM3 |  | 2.587% | 1.008% | 5.117% | 2.15% | 6.679% | -1.13% | 4.055% | 6.179% | 1.494% | -3.921% | -19.03% | 0.9002% | -5.348% | -4.692% | -14.16% | -1.099% | -39.99% | 0.5878% | 124.6%
| Z-Score (Absolute) |  | 0.27 | 0.02 | 0.67 | 0.2 | 0.92 | 0.32 | 0.5 | 0.84 | 0.09 | 0.77 | 3.19 | 0 | 1 | 0.89 | 2.41 | 0.32 | 6.54 | 0.05 | 19.78
| Analytical results for B in OREAS 242b (Indicative Value 39.8 ppm) | [21 empty] | Indicative
B | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.08 | [19 empty] | ppm
|  |  | 4A*OES | [19 empty] | 1
1
| 1 | 1 | 40 | [19 empty] | 1
| 1 | 2 | 41 | [19 empty] | 1
| 1 | 3 | 40 | [19 empty] | 16
| 1 | 4 | 39 | [19 empty] | 39.8333
| 1 | 5 | 41 | [19 empty] | 7
| 1 | 6 | 38
| Mean |  | 39.8333
| Median |  | 40
| Std Dev. |  | 1.169
| Rel.Std.Dev. |  | 2.935%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 242b (Certified Value 153 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 155.3 | 152.9227 | 145.8272 | 153 | 155 | 160 | 160 | 150 | 150 | 150 | 146.9251 | 148 | 150 | 158 | 163.6 | 145 | 164 | 152.4 | 150 | 150.538 | 1
| 1 | 2 | 155.3 | 153.4088 | 144.9926 | 147 | 158 | 160 | 160 | 150 | 150 | 160 | 153.4949 | 148 | 153 | 156 | 162.3 | 150 | 155 | 150.2 | 158 | 150.428 | 24
| 1 | 3 | 160.7 | 159.5135 | 143.5374 | 149 | 154 | 150 | 160 | 150 | 160 | 160 | 151.1527 | 143 | 155 | 150 | 167.1 | 146 | 167 | 152.5 | 151 | 505.068 | 16
| 1 | 4 | 160.5 | 152.6595 | 145.2922 | 153 | 157 | 150 | 150 | 150 | 160 | 150 | 151.7236 | 141 | 161 | 160 | 162 | 144 | 169 | 152.6 | 149 | 151.531 | 152.889
| 1 | 5 | 158.2 | 152.3099 | 143.9333 | 151 | 154 | 160 | 160 | 150 | 160 | 150 | 151.0422 | 147 | 153 | 154 | 165 | 146 | 167 | 152.3 | 149 | 149.264 | 18
| 1 | 6 | 159.4 | 154.3016 | 143.5053 | 153 | 154 | 160 | 160 | 150 | 160 | 150 | 146.7939 | 145 | 152 | 151 | 162.4 | 148 | 165 | 151.4 | 159 | 147.908
| Mean |  | 158.2333 | 154.186 | 144.5147 | 151 | 155.3333 | 156.6667 | 158.3333 | 150 | 156.6667 | 153.3333 | 150.1887 | 145.3333 | 154 | 154.8333 | 163.7333 | 146.5 | 164.5 | 151.9 | 152.6667 | 209.1228
| Median |  | 158.8 | 153.1658 | 144.463 | 152 | 154.5 | 160 | 160 | 150 | 160 | 150 | 151.0974 | 146 | 153 | 155 | 163 | 146 | 166 | 152.35 | 150.5 | 150.483
| Std Dev. |  | 2.441 | 2.6998 | 0.98667 | 2.5298 | 1.7512 | 5.164 | 4.0825 | 0 | 5.164 | 5.164 | 2.7247 | 2.8752 | 3.7947 | 3.92 | 1.9896 | 2.1679 | 4.9699 | 0.93808 | 4.5898 | 144.9882
| Rel.Std.Dev. |  | 1.543% | 1.751% | 0.6827% | 1.675% | 1.127% | 3.296% | 2.578% | 0% | 3.296% | 3.368% | 1.814% | 1.978% | 2.464% | 2.532% | 1.215% | 1.48% | 3.021% | 0.6176% | 3.006% | 69.33%
| PDM3 |  | 3.496% | 0.8483% | -5.477% | -1.236% | 1.599% | 2.471% | 3.561% | -1.89% | 2.471% | 0.2906% | -1.766% | -4.942% | 0.7267% | 1.272% | 7.093% | -4.179% | 7.594% | -0.6469% | -0.1454% | 36.78%
| Z-Score (Absolute) |  | 0.8 | 0.02 | 1.85 | 0.6 | 0.24 | 0.5 | 0.82 | 0.79 | 0.5 | 0.15 | 0.75 | 1.69 | 0.02 | 0.14 | 1.86 | 1.46 | 2.01 | 0.42 | 0.27 | 10.61
| Analytical results for Be in OREAS 242b (Certified Value 0.46 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 2
2
| 1 | 1 | 0.51 | 0.49633 | 0.34245 | 0.5 | < 5 | 0.49 | 0.46 | 0.42 | 0.45 | 0.43 | 0.57251 | 0.32 | 0.5 | < 1 | 0.51 | < 1 | 0.63 | 8.08 | 0.5 | 0.184 | 1
| 1 | 2 | 0.49 | 0.54165 | 0.29934 | 0.5 | < 5 | 0.49 | 0.46 | 0.42 | 0.5 | 0.43 | 0.5336 | 0.11 | 0.5 | < 1 | 0.49 | < 1 | 0.55 | 7.99 | 0.5 | 0.146 | 25
| 1 | 3 | 0.52 | < 0.5 | 0.36375 | 0.6 | < 5 | 0.47 | 0.45 | 0.42 | 0.47 | 0.43 | 0.58166 | 0.42 | 0.5 | < 1 | 0.48 | < 1 | 0.41 | 8.21 | 0.5 | 1.989 | 16
| 1 | 4 | 0.59 | < 0.5 | 0.36962 | 0.5 | < 5 | 0.48 | 0.45 | 0.4 | 0.47 | 0.43 | 0.5448 | 0.21 | 0.5 | < 1 | 0.46 | < 1 | 0.51 | 8.3 | 0.5 |  | 0.46117
| 1 | 5 | 0.48 | 0.54829 | 0.36909 | 0.5 | < 5 | 0.48 | 0.48 | 0.43 | 0.46 | 0.43 | 0.51876 | 0.3 | 0.5 | < 1 | 0.43 | < 1 | 0.57 | 7.82 | 0.5 | 0.143 | 19
| 1 | 6 | 0.46 | 0.56651 | 0.36297 | 0.4 | < 5 | 0.48 | 0.45 | 0.41 | 0.48 | 0.43 | 0.5803 | 0.32 | 0.5 | < 1 | 0.46 | < 1 | 0.49 | 8.03 | 0.5 | 0.143
| Mean |  | 0.50833 | 0.53819 | 0.3512 | 0.5 |  | 0.48167 | 0.45833 | 0.41667 | 0.47167 | 0.43 | 0.55527 | 0.28 | 0.5 |  | 0.47167 |  | 0.52667 | 8.0717 | 0.5 | 0.521
| Median |  | 0.5 | 0.54497 | 0.36336 | 0.5 |  | 0.48 | 0.455 | 0.42 | 0.47 | 0.43 | 0.55866 | 0.31 | 0.5 |  | 0.47 |  | 0.53 | 8.055 | 0.5 | 0.146
| Std Dev. |  | 0.04535 | 0.029825 | 0.027282 | 0.063246 |  | 0.0075277 | 0.01169 | 0.010328 | 0.017224 | 0 | 0.026578 | 0.10677 | 0 |  | 0.027869 |  | 0.075277 | 0.16916 | 0 | 0.82082
| Rel.Std.Dev. |  | 8.921% | 5.542% | 7.768% | 12.65% |  | 1.563% | 2.551% | 2.479% | 3.652% | 0% | 4.786% | 38.13% | 0% |  | 5.909% |  | 14.29% | 2.096% | 0% | 157.5%
| PDM3 |  | 10.23% | 16.7% | -23.85% | 8.42% |  | 4.444% | -0.6151% | -9.65% | 2.276% | -6.759% | 20.4% | -39.28% | 8.42% |  | 2.276% |  | 14.2% | 1650% | 8.42% | 12.97%
| Z-Score (Absolute) |  | 0.54 | 0.59 | 2.15 | NA | NA | 0.09 | 0.31 | 1.03 | 0.09 | 0.8 | 1.34 | 3.36 | NA | 0.4 | 0.09 | 0.4 | 0.85 | 129.87 | NA | 0.76
Results from laboratories 1.07, 1.08, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 242b (Certified Value 0.39 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 0.4 | 0.36765 | 0.39 | 0.58 | 0.39 | 0.44 | 0.41 | 0.37 | 0.42 | 0.36808 | 0.4 | 0.41 | 0.38 | 0.38 | 0.37 | 0.39 | < 5 | 0.4 |  |  | 1
| 1 | 2 | 0.39 | 0.37294 | 0.38 | 0.55 | 0.41 | 0.43 | 0.4 | 0.44 | 0.45 | 0.38388 | 0.37 | 0.41 | 0.36 | 0.38 | 0.38 | 0.37 | < 5 | 0.4 |  |  | 26
| 1 | 3 | 0.38 | 0.37488 | 0.37 | 0.52 | 0.39 | 0.43 | 0.42 | 0.4 | 0.42 | 0.38921 | 0.39 | 0.41 | 0.37 | 0.39 | 0.37 | 0.4 | < 5 | 0.4 | 1.504 |  | 16
| 1 | 4 | 0.4 | 0.35479 | 0.38 | 0.54 | 0.38 | 0.43 | 0.39 | 0.41 | 0.42 | 0.37129 | 0.37 | 0.41 | 0.38 | 0.42 | 0.37 | 0.38 | < 5 | 0.4 |  |  | 0.39182
| 1 | 5 | 0.41 | 0.37205 | 0.38 | 0.52 | 0.38 | 0.41 | 0.39 | 0.4 | 0.41 | 0.38627 | 0.39 | 0.41 | 0.37 | 0.38 | 0.37 | 0.39 | < 5 | 0.4 |  |  | 20
| 1 | 6 | 0.41 | 0.36519 | 0.37 | 0.51 | 0.4 | 0.43 | 0.4 | 0.38 | 0.42 | 0.38126 | 0.45 | 0.42 | 0.38 | 0.39 | 0.39 | 0.38 | < 5 | 0.4
| Mean |  | 0.39833 | 0.36792 | 0.37833 | 0.53667 | 0.39167 | 0.42833 | 0.40167 | 0.4 | 0.42333 | 0.38 | 0.395 | 0.41167 | 0.37333 | 0.39 | 0.375 | 0.385 |  | 0.4 | 1.504
| Median |  | 0.4 | 0.36985 | 0.38 | 0.53 | 0.39 | 0.43 | 0.4 | 0.4 | 0.42 | 0.38257 | 0.39 | 0.41 | 0.375 | 0.385 | 0.37 | 0.385 |  | 0.4 | 1.504
| Std Dev. |  | 0.01169 | 0.0073564 | 0.0075277 | 0.02582 | 0.01169 | 0.0098319 | 0.01169 | 0.024495 | 0.013663 | 0.0084703 | 0.029496 | 0.0040825 | 0.008165 | 0.015492 | 0.0083666 | 0.010488 |  | 6.0809e-17
| Rel.Std.Dev. |  | 2.935% | 1.999% | 1.99% | 4.811% | 2.985% | 2.295% | 2.91% | 6.124% | 3.227% | 2.229% | 7.467% | 0.9917% | 2.187% | 3.972% | 2.231% | 2.724% |  | 1.52e-14%
| PDM3 |  | 1.663% | -6.1% | -3.441% | 36.97% | -0.03826% | 9.32% | 2.514% | 2.089% | 8.044% | -3.016% | 0.8125% | 5.066% | -4.717% | -0.4636% | -4.292% | -1.74% |  | 2.089% | 283.9%
| Z-Score (Absolute) |  | 0.06 | 1.11 | 0.71 | 5.39 | 0.19 | 1.22 | 0.19 | 0.13 | 1.03 | 0.64 | 0.06 | 0.58 | 0.9 | 0.26 | 0.83 | 0.45 | 81.05 | NA | 42.67
| Analytical results for Ca in OREAS 242b (Certified Value 6.31 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 6.6554 | 6.2229 | 6.4964 | 5.91 | 6.3 | 6.41 | 6.5 | 6.32 | 5.99 | 6.03 | 6.4836 | 6.7748 | 6.36 | 6.23 | 6.0285 | 5.81 | 6.99 | 6.35 | 6.2 | 6.5165 | 1
| 1 | 2 | 6.5621 | 6.1466 | 6.4924 | 5.98 | 6.21 | 6.29 | 6.54 | 6.2 | 6.46 | 6.07 | 6.5459 | 6.8886 | 6.4 | 6.49 | 5.9832 | 6.05 | 6.5 | 6.2 | 6.14 | 6.4825 | #N/A
| 1 | 3 | 6.6001 | 6.3895 | 6.5018 | 5.88 | 6.32 | 6.31 | 6.49 | 6.1 | 6.28 | 6.03 | 6.5105 | 6.8321 | 6.29 | 6.05 | 6.1385 | 5.94 | 6.1 | 6.31 | 6.15 | 1.7201 | 16
| 1 | 4 | 6.6394 | 6.2486 | 6.5123 | 5.93 | 6.32 | 6.31 | 6.48 | 6.02 | 6.22 | 5.95 | 6.5401 | 6.7713 | 6.49 | 6.52 | 6.0181 | 6 | 6.6 | 6.3 | 6.15 | 6.5214 | 6.3111
| 1 | 5 | 6.7477 | 6.1757 | 6.5067 | 6.06 | 6.33 | 6.32 | 6.51 | 5.95 | 6.22 | 5.96 | 6.4668 | 6.7238 | 6.46 | 6.33 | 6.0435 | 5.95 | 6.14 | 6.26 | 6.16 | 6.4245 | 21
| 1 | 6 | 6.6488 | 6.295 | 6.5116 | 6.01 | 6.22 | 6.39 | 6.56 | 6.21 | 6.37 | 5.85 | 6.5713 | 6.8391 | 6.3 | 6.09 | 5.9844 | 6.04 | 6.34 | 6.25 | 6.17 | 6.4799
| Mean |  | 6.6423 | 6.2464 | 6.5035 | 5.9617 | 6.2833 | 6.3383 | 6.5133 | 6.1333 | 6.2567 | 5.9817 | 6.5197 | 6.805 | 6.3833 | 6.285 | 6.0327 | 5.965 | 6.445 | 6.2783 | 6.1617 | 5.6908
| Median |  | 6.6441 | 6.2357 | 6.5043 | 5.955 | 6.31 | 6.315 | 6.505 | 6.15 | 6.25 | 5.995 | 6.5253 | 6.8034 | 6.38 | 6.28 | 6.0233 | 5.975 | 6.42 | 6.28 | 6.155 | 6.4812
| Std Dev. |  | 0.062521 | 0.087579 | 0.0081153 | 0.067355 | 0.053914 | 0.04916 | 0.030768 | 0.13619 | 0.16033 | 0.079099 | 0.039879 | 0.059207 | 0.082138 | 0.19776 | 0.057139 | 0.088261 | 0.3308 | 0.052694 | 0.02137 | 1.9456
| Rel.Std.Dev. |  | 0.9413% | 1.402% | 0.1248% | 1.13% | 0.858% | 0.7756% | 0.4724% | 2.22% | 2.563% | 1.322% | 0.6117% | 0.8701% | 1.287% | 3.147% | 0.9472% | 1.48% | 5.133% | 0.8393% | 0.3468% | 34.19%
| PDM3 |  | 5.248% | -1.025% | 3.05% | -5.536% | -0.4393% | 0.4322% | 3.205% | -2.816% | -0.8619% | -5.219% | 3.306% | 7.826% | 1.145% | -0.4129% | -4.411% | -5.483% | 2.122% | -0.5185% | -2.367% | -9.828%
| Z-Score (Absolute) |  | 1.26 | 0.12 | 0.78 | 1.11 | 0.01 | 0.2 | 0.81 | 0.51 | 0.08 | 1.04 | 0.83 | 1.83 | 0.36 | 0.01 | 0.87 | 1.1 | 0.57 | 0.01 | 0.42 | 2.06
| Analytical results for Cd in OREAS 242b (Certified Value 0.4 ppm) | [21 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 0.42 | 0.37366 | 0.5 | 0.48 | 0.41 | 0.41 | 0.4 | 0.39 | 0.44 | 0.38161 | 0.37 | 0.37 | 0.41 | 0.43 | 0.35 | 0.27 | < 1 | < 0.5 | 2.198 |  | 1
| 1 | 2 | 0.4 | 0.33883 | 0.4 | 0.43 | 0.42 | 0.43 | 0.4 | 0.43 | 0.41 | 0.33078 | 0.36 | 0.43 | 0.4 | 0.46 | 0.37 | 0.3 | < 1 | < 0.5 | 2.08 |  | 27
| 1 | 3 | 0.44 | 0.3594 | 0.3 | 0.42 | 0.4 | 0.45 | 0.38 | 0.42 | 0.42 | 0.37036 | 0.37 | 0.42 | 0.41 | 0.41 | 0.36 | 0.32 | < 1 | < 0.5 | 0.953 |  | 16
| 1 | 4 | 0.39 | 0.35974 | 0.6 | 0.41 | 0.37 | 0.43 | 0.43 | 0.4 | 0.41 | 0.36188 | 0.36 | 0.41 | 0.36 | 0.36 | 0.37 | 0.32 | < 1 | < 0.5 | 2.11 |  | 0.39693
| 1 | 5 | 0.36 | 0.39236 | 0.4 | 0.41 | 0.4 | 0.43 | 0.36 | 0.39 | 0.42 | 0.37879 | 0.38 | 0.4 | 0.44 | 0.41 | 0.35 | 0.3 | < 1 | < 0.5 | 2.123 |  | 22
| 1 | 6 | 0.41 | 0.37147 | 0.4 | 0.44 | 0.42 | 0.44 | 0.42 | 0.4 | 0.41 | 0.33358 | 0.4 | 0.41 | 0.32 | 0.39 | 0.36 | 0.28 | < 1 | < 0.5 | 2.016
| Mean |  | 0.40333 | 0.36591 | 0.43333 | 0.43167 | 0.40333 | 0.43167 | 0.39833 | 0.405 | 0.41833 | 0.3595 | 0.37333 | 0.40667 | 0.39 | 0.41 | 0.36 | 0.29833 |  |  | 1.9133
| Median |  | 0.405 | 0.36561 | 0.4 | 0.425 | 0.405 | 0.43 | 0.4 | 0.4 | 0.415 | 0.36612 | 0.37 | 0.41 | 0.405 | 0.41 | 0.36 | 0.3 |  |  | 2.095
| Std Dev. |  | 0.027325 | 0.017914 | 0.10328 | 0.026394 | 0.018619 | 0.013292 | 0.025626 | 0.016432 | 0.01169 | 0.022279 | 0.015055 | 0.020656 | 0.042895 | 0.034059 | 0.0089443 | 0.020412 |  |  | 0.47418
| Rel.Std.Dev. |  | 6.775% | 4.896% | 23.83% | 6.115% | 4.616% | 3.079% | 6.433% | 4.057% | 2.795% | 6.197% | 4.033% | 5.079% | 11% | 8.307% | 2.485% | 6.842% |  |  | 24.78%
| PDM3 |  | 1.612% | -7.816% | 9.17% | 8.75% | 1.612% | 8.75% | 0.3525% | 2.032% | 5.391% | -9.43% | -5.946% | 2.452% | -1.747% | 3.292% | -9.305% | -24.84% |  |  | 382%
| Z-Score (Absolute) |  | 0 | 0.89 | NA | 0.67 | 0 | 0.67 | 0.12 | 0.04 | 0.36 | 1.04 | 0.71 | 0.08 | 0.32 | 0.16 | 1.03 | 2.5 | 2.3 | 3.65 | 35.94
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 242b (Certified Value 17 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 1
2
| 1 | 1 | 16.84 | 17.3505 | 16.6054 | 17 | 18.15 | 16.95 | 18.05 | 16.4 | 19.05 | 17.6 | 16.51 | 16.63 | 16.85 | 16.2 | 17.2 | 45.14 |  |  |  |  | 1
| 1 | 2 | 16.7 | 17.0756 | 16.5688 | 16.5 | 18.35 | 17 | 17.75 | 17.95 | 20.1 | 17.4 | 15.93 | 16.43 | 17.08 | 15.7 | 17.9 | 44.72 |  |  |  |  | 28
| 1 | 3 | 16.89 | 18.0813 | 16.5555 | 16.6 | 17.8 | 17.2 | 18.3 | 17.5 | 17.9 | 17.2 | 15.81 | 16.68 | 17.06 | 16.1 | 17.8 | 84.797 |  |  |  |  | 16
| 1 | 4 | 16.93 | 17.5201 | 16.4232 | 16.7 | 17.5 | 16.95 | 17.5 | 17.55 | 18.8 | 17 | 16.17 | 16.12 | 16.78 | 16.2 | 17.5 | 44.577 |  |  |  |  | 16.9627
| 1 | 5 | 16.97 | 16.9461 | 16.791 | 16.8 | 17.7 | 18.45 | 17.1 | 17 | 17.95 | 17.2 | 15.79 | 16.6 | 16.88 | 16.1 | 16.8 | 44.38 |  |  |  |  | 23
| 1 | 6 | 16.91 | 17.1938 | 16.6208 | 17 | 18 | 16.9 | 17.5 | 16.1 | 19.3 | 17.4 | 15.74 | 16.44 | 17.12 | 16.6 | 17.2 | 44.04
| Mean |  | 16.8733 | 17.3612 | 16.5941 | 16.7667 | 17.9167 | 17.2417 | 17.7 | 17.0833 | 18.85 | 17.3 | 15.9917 | 16.4833 | 16.9617 | 16.15 | 17.4 | 51.2757
| Median |  | 16.9 | 17.2722 | 16.5871 | 16.75 | 17.9 | 16.975 | 17.625 | 17.25 | 18.925 | 17.3 | 15.87 | 16.52 | 16.97 | 16.15 | 17.35 | 44.6485
| Std Dev. |  | 0.095219 | 0.40641 | 0.1191 | 0.20656 | 0.31091 | 0.60118 | 0.43012 | 0.7188 | 0.83905 | 0.20976 | 0.29708 | 0.2052 | 0.14204 | 0.2881 | 0.41473 | 16.4261
| Rel.Std.Dev. |  | 0.5643% | 2.341% | 0.7177% | 1.232% | 1.735% | 3.487% | 2.43% | 4.208% | 4.451% | 1.212% | 1.858% | 1.245% | 0.8374% | 1.784% | 2.383% | 32.03%
| PDM3 |  | -0.5271% | 2.349% | -2.173% | -1.156% | 5.624% | 1.644% | 4.346% | 0.7109% | 11.13% | 1.988% | -5.725% | -2.826% | -0.006326% | -4.791% | 2.578% | 202.3%
| Z-Score (Absolute) |  | 0.42 | 0.29 | 0.82 | 0.57 | 1.09 | 0.11 | 0.78 | 0.11 | 2.44 | 0.2 | 1.69 | 0.98 | 0.29 | 1.46 | 0.34 | 49.29
| Analytical results for Co in OREAS 242b (Certified Value 39.6 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 41.3 | 39.6433 | 38.297 | 37 | 39 | 40.4 | 40.5 | 41.4 | 38.6 | 38.1 | 38.2929 | 42.6 | 39 | 39.6 | 37.9 | 39.2 | 40.7 | 39.86 | 40 | 39.382 | 1
| 1 | 2 | 40.8 | 38.6498 | 38.4221 | 36 | 38 | 39.1 | 39.8 | 40.3 | 42 | 40.5 | 40.2418 | 41.8 | 40 | 38.7 | 37.9 | 40.6 | 40.2 | 39.46 | 40 | 40.237 | 14
| 1 | 3 | 42.2 | 40.8512 | 38.0227 | 36 | 38 | 38.9 | 40.1 | 41.5 | 41 | 37.7 | 39.0709 | 41.6 | 39 | 38 | 39.1 | 40.5 | 40.6 | 39.73 | 41 | 18.029 | 16
| 1 | 4 | 41.6 | 39.6594 | 38.5899 | 38 | 39 | 38.1 | 39.5 | 40.8 | 42.2 | 39.3 | 40.2799 | 41 | 39 | 39.7 | 38.2 | 40.5 | 41.9 | 39.42 | 41 | 40.021 | 39.6156
| 1 | 5 | 40.9 | 39.064 | 37.9626 | 37 | 39 | 39.8 | 42.2 | 39 | 40.9 | 37.5 | 39.2527 | 41.4 | 39 | 39.4 | 38.6 | 40.4 | 41.1 | 38.85 | 41 | 39.526 | 24
| 1 | 6 | 41.6 | 39.8149 | 38.2181 | 37 | 39 | 39.7 | 40.9 | 39.9 | 39.5 | 38.2 | 37.8933 | 41.5 | 39 | 36.8 | 37.9 | 40.5 | 41.4 | 39.42 | 40 | 38.753
| Mean |  | 41.4 | 39.6138 | 38.2521 | 36.8333 | 38.6667 | 39.3333 | 40.5 | 40.4833 | 40.7 | 38.55 | 39.1719 | 41.65 | 39.1667 | 38.7 | 38.2667 | 40.2833 | 40.9833 | 39.4567 | 40.5 | 35.9913
| Median |  | 41.45 | 39.6514 | 38.2576 | 37 | 39 | 39.4 | 40.3 | 40.55 | 40.95 | 38.15 | 39.1618 | 41.55 | 39 | 39.05 | 38.05 | 40.5 | 40.9 | 39.44 | 40.5 | 39.454
| Std Dev. |  | 0.51769 | 0.74915 | 0.23774 | 0.75277 | 0.5164 | 0.80664 | 0.96954 | 0.95376 | 1.4085 | 1.1415 | 0.97945 | 0.53572 | 0.40825 | 1.1314 | 0.49261 | 0.53448 | 0.61128 | 0.34875 | 0.54772 | 8.8151
| Rel.Std.Dev. |  | 1.25% | 1.891% | 0.6215% | 2.044% | 1.336% | 2.051% | 2.394% | 2.356% | 3.461% | 2.961% | 2.5% | 1.286% | 1.042% | 2.923% | 1.287% | 1.327% | 1.492% | 0.8839% | 1.352% | 24.49%
| PDM3 |  | 4.504% | -0.004572% | -3.442% | -7.023% | -2.395% | -0.7125% | 2.233% | 2.19% | 2.737% | -2.69% | -1.12% | 5.135% | -1.133% | -2.311% | -3.405% | 1.686% | 3.453% | -0.4011% | 2.233% | -9.149%
| Z-Score (Absolute) |  | 1.23 | 0.13 | 0.7 | 1.58 | 0.45 | 0.04 | 0.68 | 0.67 | 0.8 | 0.52 | 0.14 | 1.39 | 0.14 | 0.43 | 0.69 | 0.55 | 0.98 | 0.04 | 0.68 | 2.09
| Analytical results for Cr in OREAS 242b (Certified Value 107 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 131 | 104.8081 | 64 | 122 | 103 | 91 | 106 | 102 | 106 | 97.3282 | 114.8 | 113 | 103 | 100 | 102 | 126 | 110.7 | 130 | 130.247 |  | 1
| 1 | 2 | 131 | 104.4581 | 62 | 118 | 103 | 84 | 104 | 104 | 106 | 98.4585 | 117.4 | 113 | 109 | 97 | 109 | 124 | 125.1 | 130 | 131.979 |  | 30
| 1 | 3 | 127 | 112.8088 | 78 | 120 | 102 | 96 | 105 | 105 | 106 | 94.6588 | 115.6 | 101 | 107 | 102 | 106 | 115 | 112.4 | 120 | 136.578 |  | 16
| 1 | 4 | 132 | 105.9539 | 115 | 125 | 105 | 93 | 104 | 105 | 103 | 100.5648 | 115.3 | 99 | 112 | 105 | 108 | 128 | 111.6 | 130 | 131.468 |  | 106.9892
| 1 | 5 | 130 | 106.1581 | 76 | 122 | 104 | 99 | 100 | 107 | 104 | 100.7927 | 114.7 | 114 | 112 | 111 | 105 | 112 | 108.8 | 130 | 131.007 |  | 25
| 1 | 6 | 128 | 108.8169 | 71 | 123 | 102 | 89 | 104 | 107 | 99 | 95.1344 | 116.5 | 106 | 99 | 105 | 105 | 115 | 115.3 | 130 | 128.865
| Mean |  | 129.8333 | 107.1673 | 77.6667 | 121.6667 | 103.1667 | 92 | 103.8333 | 105 | 104 | 97.8229 | 115.7167 | 107.6667 | 107 | 103.3333 | 105.8333 | 120 | 113.9833 | 128.3333 | 131.6907
| Median |  | 130.5 | 106.056 | 73.5 | 122 | 103 | 92 | 104 | 105 | 105 | 97.8934 | 115.45 | 109.5 | 108 | 103.5 | 105.5 | 119.5 | 112 | 130 | 131.2375
| Std Dev. |  | 1.9408 | 3.1605 | 19.3563 | 2.4221 | 1.169 | 5.2915 | 2.0412 | 1.8974 | 2.7568 | 2.6163 | 1.0496 | 6.6232 | 5.1769 | 4.8442 | 2.4833 | 6.7823 | 5.8506 | 4.0825 | 2.6289
| Rel.Std.Dev. |  | 1.495% | 2.949% | 24.92% | 1.991% | 1.133% | 5.752% | 1.966% | 1.807% | 2.651% | 2.675% | 0.907% | 6.152% | 4.838% | 4.688% | 2.346% | 5.652% | 5.133% | 3.181% | 1.996%
| PDM3 |  | 21.35% | 0.1664% | -27.41% | 13.72% | -3.573% | -14.01% | -2.95% | -1.859% | -2.794% | -8.568% | 8.157% | 0.6332% | 0.01006% | -3.417% | -1.08% | 12.16% | 6.537% | 19.95% | 23.09%
| Z-Score (Absolute) |  | 2.2 | 0.02 | 2.83 | 1.42 | 0.37 | 1.45 | 0.31 | 0.19 | 0.29 | 0.89 | 0.84 | 0.06 | 0 | 0.35 | 0.11 | 1.26 | 0.67 | 2.06 | 2.38
| Analytical results for Cs in OREAS 242b (Certified Value 1.08 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 1.11 | 1.0863 | 1.08 | 1.09 | 1.1 | 1.1 | 1.02 | 1.19 | 1.0729 | 1 | 1.1 | 1.1 | 1.13 | 1 | 1.2 | 1.1 |  |  |  |  | 1
| 1 | 2 | 1.09 | 1.0403 | 1.07 | 1.11 | 1.12 | 1.04 | 1.17 | 1.28 | 1.0944 | 1 | 1.1 | 1 | 1.11 | 1.1 | 1.1 | 1 |  |  |  |  | 31
| 1 | 3 | 1.14 | 1.0921 | 1.07 | 1.05 | 1.12 | 1.06 | 1.11 | 1.15 | 1.1123 | 1 | 1.1 | 1 | 1.18 | 1 | 1.2 | 1.1 |  |  |  |  | 16
| 1 | 4 | 1.1 | 1.0864 | 1.11 | 1.04 | 1.08 | 1.04 | 1.09 | 1.18 | 1.127 | 1 | 1.1 | 1.1 | 1.14 | 1 | 1.2 | 1 |  |  |  |  | 1.0795
| 1 | 5 | 1.11 | 1.0042 | 1.03 | 1.06 | 1.16 | 1.01 | 1.02 | 1.15 | 1.1034 | 1 | 1.1 | 1 | 1.12 | 1.1 | 1.2 | 1.1 |  |  |  |  | 26
| 1 | 6 | 1.12 | 1.0809 | 1.09 | 1.07 | 1.12 | 1.07 | 1.02 | 1.2 | 1.0902 | 1 | 1.1 | 1 | 1.12 | 1.1 | 1.2 | 1
| Mean |  | 1.1117 | 1.065 | 1.075 | 1.07 | 1.1167 | 1.0533 | 1.0717 | 1.1917 | 1.1 | 1 | 1.1 | 1.0333 | 1.1333 | 1.05 | 1.1833 | 1.05
| Median |  | 1.11 | 1.0836 | 1.075 | 1.065 | 1.12 | 1.05 | 1.055 | 1.185 | 1.0989 | 1 | 1.1 | 1 | 1.125 | 1.05 | 1.2 | 1.05
| Std Dev. |  | 0.017224 | 0.03523 | 0.026646 | 0.026077 | 0.026583 | 0.030768 | 0.062423 | 0.047924 | 0.018726 | 0 | 2.4324e-16 | 0.05164 | 0.025033 | 0.054772 | 0.040825 | 0.054772
| Rel.Std.Dev. |  | 1.549% | 3.308% | 2.479% | 2.437% | 2.381% | 2.921% | 5.825% | 4.022% | 1.702% | 0% | 2.211e-14% | 4.997% | 2.209% | 5.216% | 3.45% | 5.216%
| PDM3 |  | 2.984% | -1.335% | -0.4127% | -0.8759% | 3.447% | -2.42% | -0.7215% | 10.4% | 1.907% | -7.361% | 1.903% | -4.273% | 4.991% | -2.729% | 9.623% | -2.729%
| Z-Score (Absolute) |  | 0.99 | 0.27 | 0 | 0.13 | 1.12 | 0.58 | 0.09 | 3.15 | 0.68 | NA | 0.67 | 1.12 | 1.57 | 0.67 | 2.92 | 0.67
Results from laboratory 1.15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 242b (Certified Value 173 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 175 | 171.2357 | 177.067 | 174 | 181 | 175 | 172.5 | 181 | 168.5 | 172 | 172.2299 | 153.9 | 178 | 178.4 | 165.3 | 155.8 | 201.1 | 156.4 | 162 | 179.169 | 1
| 1 | 2 | 170.4 | 170.1415 | 177.2455 | 174 | 190 | 174.5 | 174.5 | 177 | 179.5 | 178 | 169.2295 | 156.1 | 180 | 176 | 161.2 | 160.3 | 196.1 | 152.9 | 166 | 178.359 | 16
| 1 | 3 | 175.4 | 179.3631 | 174.6205 | 177 | 181 | 168.5 | 174.5 | 178 | 174.5 | 173.5 | 169.6106 | 156.3 | 178 | 168.3 | 169.4 | 159 | 199.7 | 154.6 | 164 | 33.558 | 16
| 1 | 4 | 174.1 | 173.213 | 175.009 | 175 | 177 | 172 | 171.5 | 176 | 178.5 | 173.5 | 168.6501 | 154.9 | 176 | 181.4 | 165.4 | 161.6 | 208.8 | 158.4 | 160 | 178.586 | 173.3828
| 1 | 5 | 173.6 | 169.4086 | 174.4735 | 178 | 176 | 172.5 | 174.5 | 172 | 174.5 | 169 | 172.5312 | 152.8 | 176 | 171.5 | 163 | 160.3 | 198.3 | 148.1 | 160 | 175.747 | 27
| 1 | 6 | 174.5 | 175.4669 | 174.5785 | 178 | 178 | 173.5 | 175.5 | 180 | 177 | 170.5 | 168.466 | 158.7 | 176 | 168.9 | 161.2 | 162.5 | 200.9 | 158 | 160 | 174.481
| Mean |  | 173.8333 | 173.1381 | 175.499 | 176 | 180.5 | 172.6667 | 173.8333 | 177.3333 | 175.4167 | 172.75 | 170.1196 | 155.45 | 177.3333 | 174.0833 | 164.25 | 159.9167 | 200.8167 | 154.7333 | 162 | 153.3167
| Median |  | 174.3 | 172.2243 | 174.8148 | 176 | 179.5 | 173 | 174.5 | 177.5 | 175.75 | 172.75 | 169.42 | 155.5 | 177 | 173.75 | 164.15 | 160.3 | 200.3 | 155.5 | 161 | 177.053
| Std Dev. |  | 1.7985 | 3.76 | 1.2977 | 1.8974 | 5.0892 | 2.3381 | 1.5055 | 3.2042 | 3.955 | 3.1105 | 1.8008 | 2.0705 | 1.633 | 5.3447 | 3.1342 | 2.349 | 4.3278 | 3.8542 | 2.5298 | 58.6978
| Rel.Std.Dev. |  | 1.035% | 2.172% | 0.7394% | 1.078% | 2.82% | 1.354% | 0.8661% | 1.807% | 2.255% | 1.801% | 1.059% | 1.332% | 0.9209% | 3.07% | 1.908% | 1.469% | 2.155% | 2.491% | 1.562% | 38.29%
| PDM3 |  | 0.2598% | -0.1411% | 1.221% | 1.509% | 4.105% | -0.413% | 0.2598% | 2.278% | 1.173% | -0.365% | -1.882% | -10.34% | 2.278% | 0.404% | -5.267% | -7.767% | 15.82% | -10.76% | -6.565% | -11.57%
| Z-Score (Absolute) |  | 0.06 | 0.06 | 0.38 | 0.47 | 1.31 | 0.15 | 0.06 | 0.72 | 0.36 | 0.14 | 0.63 | 3.37 | 0.72 | 0.11 | 1.73 | 2.54 | 5.11 | 3.51 | 2.15 | 3.77
| Analytical results for Dy in OREAS 242b (Certified Value 3.83 ppm) | [21 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 3.88 | 4.0211 | 4.3683 | 4.1 | 3.9 | 3.76 | 3.6 | 3.58 | 3.8 | 4.194 | [10 empty] | 1
| 1 | 2 | 4 | 3.8646 | 4.348 | 3.8 | 3.8 | 3.83 | 3.7 | 3.6 | 3.8 | 4.128 | [10 empty] | 4
| 1 | 3 | 3.87 | 3.8649 | 4.386 | 3.9 | 3.8 | 3.89 | 3.6 | 3.74 | 3.8 | 4.095 | [10 empty] | 16
| 1 | 4 | 4.07 | 3.7514 | 4.3777 | 3.8 | 3.8 | 3.84 | 3.6 | 3.63 | 3.8 | 4.084 | [10 empty] | 3.8325
| 1 | 5 | 3.9 | 3.8231 | 4.3609 | 4 | 3.6 | 3.79 | 3.6 | 3.68 | 3.85 | 4.173 | [10 empty] | 28
| 1 | 6 | 3.95 | 3.7546 | 4.3133 | 3.8 | 3.8 | 3.76 | 3.6 | 3.76 | 3.75 | 4.069
| Mean |  | 3.945 | 3.8466 | 4.359 | 3.9 | 3.7833 | 3.8117 | 3.6167 | 3.665 | 3.8 | 4.1238
| Median |  | 3.925 | 3.8438 | 4.3646 | 3.85 | 3.8 | 3.81 | 3.6 | 3.655 | 3.8 | 4.1115
| Std Dev. |  | 0.078166 | 0.099167 | 0.025991 | 0.12649 | 0.098319 | 0.051153 | 0.040825 | 0.074229 | 0.031623 | 0.050563
| Rel.Std.Dev. |  | 1.981% | 2.578% | 0.5962% | 3.243% | 2.599% | 1.342% | 1.129% | 2.025% | 0.8322% | 1.226%
| PDM3 |  | 2.937% | 0.3693% | 13.74% | 1.762% | -1.282% | -0.5425% | -5.631% | -4.369% | -0.8469% | 7.603%
| Z-Score (Absolute) |  | 0.84 | 0.13 | 3.83 | 0.51 | 0.33 | 0.13 | 1.53 | 1.19 | 0.21 | 2.13
| Analytical results for Er in OREAS 242b (Certified Value 2.3 ppm) | [21 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.45 | 2.3776 | 2.8158 | 2.6 | 2.4 | 2.16 | 2.1 | 2.3 | 2.15 | [11 empty] | 1
| 1 | 2 | 2.37 | 2.2941 | 2.7617 | 2.4 | 2.4 | 2.21 | 2.2 | 2.23 | 2.2 | [11 empty] | 5
| 1 | 3 | 2.41 | 2.3693 | 2.7267 | 2.3 | 2.4 | 2.27 | 2.1 | 2.32 | 2.15 | [11 empty] | 16
| 1 | 4 | 2.43 | 2.2781 | 2.8428 | 2.4 | 2.5 | 2.22 | 2.1 | 2.19 | 2.15 | [11 empty] | 2.2975
| 1 | 5 | 2.41 | 2.2491 | 2.8144 | 2.6 | 2.4 | 2.23 | 2.1 | 2.33 | 2.2 | [11 empty] | 29
| 1 | 6 | 2.43 | 2.2313 | 2.7165 | 2.4 | 2.3 | 2.27 | 2.2 | 2.3 | 2.2
| Mean |  | 2.4167 | 2.2999 | 2.7797 | 2.45 | 2.4 | 2.2267 | 2.1333 | 2.2783 | 2.175
| Median |  | 2.42 | 2.2861 | 2.7881 | 2.4 | 2.4 | 2.225 | 2.1 | 2.3 | 2.175
| Std Dev. |  | 0.027325 | 0.061057 | 0.052168 | 0.12247 | 0.063246 | 0.041312 | 0.05164 | 0.055648 | 0.027386
| Rel.Std.Dev. |  | 1.131% | 2.655% | 1.877% | 4.999% | 2.635% | 1.855% | 2.421% | 2.442% | 1.259%
| PDM3 |  | 5.187% | 0.105% | 20.99% | 6.638% | 4.462% | -3.083% | -7.145% | -0.8337% | -5.331%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.77 | 0.87 | 0.58 | 0.42 | 0.96 | 0.12 | 0.72
| Analytical results for Eu in OREAS 242b (Certified Value 0.9 ppm) | [21 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 0.9 | 0.89806 | 1.0621 | 0.92 | 0.9 | 0.83 | 0.8 | 0.79 | 0.95 | 1.36 | [10 empty] | 1
| 1 | 2 | 0.89 | 0.87252 | 1.0358 | 0.92 | 0.9 | 0.82 | 0.8 | 0.78 | 0.95 | 1.36 | [10 empty] | 35
| 1 | 3 | 0.88 | 0.90273 | 1.0467 | 0.9 | 0.8 | 0.85 | 0.8 | 0.82 | 0.95 | 1.672 | [10 empty] | 16
| 1 | 4 | 0.93 | 0.8868 | 1.0471 | 0.84 | 0.9 | 0.83 | 0.8 | 0.79 | 0.95 | 1.356 | [10 empty] | 0.90098
| 1 | 5 | 0.89 | 0.89249 | 1.054 | 0.89 | 0.8 | 0.83 | 0.8 | 0.81 | 0.95 | 1.313 | [10 empty] | 30
| 1 | 6 | 0.93 | 0.86519 | 0.99756 | 0.88 | 0.8 | 0.83 | 0.8 | 0.83 | 0.95 | 1.28
| Mean |  | 0.90333 | 0.8863 | 1.0405 | 0.89167 | 0.85 | 0.83167 | 0.8 | 0.80333 | 0.95 | 1.3902
| Median |  | 0.895 | 0.88965 | 1.0469 | 0.895 | 0.85 | 0.83 | 0.8 | 0.8 | 0.95 | 1.358
| Std Dev. |  | 0.021602 | 0.014711 | 0.022781 | 0.029944 | 0.054772 | 0.0098319 | 1.2162e-16 | 0.019664 | 1.2162e-16 | 0.14178
| Rel.Std.Dev. |  | 2.391% | 1.66% | 2.189% | 3.358% | 6.444% | 1.182% | 1.52e-14% | 2.448% | 1.28e-14% | 10.2%
| PDM3 |  | 0.2616% | -1.629% | 15.49% | -1.033% | -5.658% | -7.693% | -11.21% | -10.84% | 5.441% | 54.3%
| Z-Score (Absolute) |  | 0.07 | 0.13 | 1.63 | 0.07 | NA | 0.75 | NA | 1.07 | 0.6 | 5.61
Results from laboratories 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 242b (Certified Value 7.5 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.67 | 7.2306 | 7.5607 | 7.41 | 7.93 | 7.37 | 7.46 | 7.7 | 7.13 | 7.29 | 8.1291 | 7.6768 | 7.88 | 7.23 | 7.0228 | 6.81 | 7.56 | 7.47 | 7.64 | 8.1137 | 1
| 1 | 2 | 7.58 | 7.217 | 7.5531 | 7.21 | 7.41 | 7.29 | 7.66 | 7.6 | 7.8 | 7.33 | 8.023 | 7.7363 | 7.71 | 7.4 | 7.0324 | 7.03 | 7.42 | 7.4 | 7.55 | 8.0675 | #N/A
| 1 | 3 | 7.6 | 7.6201 | 7.5615 | 7.32 | 7.45 | 7.23 | 7.56 | 7.49 | 7.5 | 7.34 | 7.9076 | 7.6956 | 7.88 | 7.03 | 7.2138 | 6.88 | 7.46 | 7.52 | 7.65 | 4.6717 | 16
| 1 | 4 | 7.75 | 7.3093 | 7.5614 | 7.37 | 7.58 | 7.28 | 7.41 | 7.43 | 7.48 | 7.31 | 8.1464 | 7.6555 | 7.87 | 7.55 | 7.0314 | 7.01 | 7.68 | 7.5 | 7.6 | 8.1236 | 7.5002
| 1 | 5 | 7.74 | 7.1249 | 7.5627 | 7.45 | 7.66 | 7.31 | 7.44 | 7.35 | 7.51 | 7.21 | 7.9266 | 7.621 | 7.77 | 7.31 | 7.0718 | 6.97 | 7.46 | 7.44 | 7.65 | 8.0161 | 31
| 1 | 6 | 7.63 | 7.3217 | 7.5597 | 7.3 | 7.56 | 7.36 | 7.61 | 7.53 | 7.51 | 7.19 | 7.898 | 7.7925 | 7.63 | 6.95 | 6.957 | 7.03 | 7.55 | 7.43 | 7.57 | 8.0627
| Mean |  | 7.6617 | 7.3039 | 7.5598 | 7.3433 | 7.5983 | 7.3067 | 7.5233 | 7.5167 | 7.4883 | 7.2783 | 8.0051 | 7.6963 | 7.79 | 7.245 | 7.0549 | 6.955 | 7.5217 | 7.46 | 7.61 | 7.5092
| Median |  | 7.65 | 7.2699 | 7.561 | 7.345 | 7.57 | 7.3 | 7.51 | 7.51 | 7.505 | 7.3 | 7.9748 | 7.6862 | 7.82 | 7.27 | 7.0319 | 6.99 | 7.505 | 7.455 | 7.62 | 8.0651
| Std Dev. |  | 0.071391 | 0.17045 | 0.0034504 | 0.085713 | 0.18606 | 0.052409 | 0.10132 | 0.12388 | 0.21292 | 0.063377 | 0.11206 | 0.060942 | 0.10488 | 0.22572 | 0.086265 | 0.090719 | 0.095167 | 0.045166 | 0.043359 | 1.3906
| Rel.Std.Dev. |  | 0.9318% | 2.334% | 0.04564% | 1.167% | 2.449% | 0.7173% | 1.347% | 1.648% | 2.843% | 0.8708% | 1.4% | 0.7918% | 1.346% | 3.116% | 1.223% | 1.304% | 1.265% | 0.6054% | 0.5698% | 18.52%
| PDM3 |  | 2.153% | -2.616% | 0.7955% | -2.091% | 1.309% | -2.58% | 0.3087% | 0.2199% | -0.1579% | -2.958% | 6.732% | 2.615% | 3.864% | -3.402% | -5.937% | -7.269% | 0.2865% | -0.5357% | 1.464% | 0.1207%
| Z-Score (Absolute) |  | 0.63 | 0.89 | 0.2 | 0.72 | 0.36 | 0.87 | 0.04 | 0.02 | 0.1 | 0.99 | 2.09 | 0.78 | 1.17 | 1.14 | 1.94 | 2.36 | 0.04 | 0.22 | 0.41 | 0.02
| Analytical results for Ga in OREAS 242b (Certified Value 15.8 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 15.95 | 15.5296 | 17.417 | 16 | 16.1 | 15.75 | 15.4 | 15.95 | 15.55 | 15.15 | 14.9848 | 17.25 | 16.8 | 15.43 | 16.01 | 14.36 | 15 | 16.4 |  |  | 1
| 1 | 2 | 16.39 | 15.1873 | 17.499 | 15.8 | 16.1 | 15.9 | 15.35 | 15.75 | 17 | 16.3 | 15.835 | 17.04 | 17.2 | 14.33 | 16.39 | 14.65 | 14.6 | 16.2 |  |  | 36
| 1 | 3 | 16.52 | 16.2852 | 17.461 | 16 | 15.8 | 15.25 | 15.5 | 15.95 | 16.25 | 15.45 | 15.1994 | 16.81 | 17.2 | 14.78 | 16.45 | 14.5 | 15.5 | 16.4 |  |  | 16
| 1 | 4 | 16.51 | 15.7988 | 17.435 | 15.8 | 16.3 | 15.35 | 15 | 15.55 | 16.35 | 15.75 | 16.1666 | 16.56 | 17.1 | 14.75 | 15.97 | 14.77 | 15.5 | 16.4 |  |  | 15.8051
| 1 | 5 | 16.4 | 15.2289 | 17.442 | 16.3 | 16.2 | 15.55 | 15.9 | 15.25 | 15.8 | 15.15 | 15.7324 | 16.7 | 16.9 | 14.4 | 16.11 | 14.51 | 15.1 | 16.4 |  |  | 32
| 1 | 6 | 16.43 | 15.5003 | 17.444 | 16 | 15.7 | 15.55 | 15.7 | 16 | 15.65 | 15.3 | 15.1156 | 16.94 | 16.8 | 14.46 | 15.89 | 14.65 | 15.5 | 16.2
| Mean |  | 16.3667 | 15.5883 | 17.4497 | 15.9833 | 16.0333 | 15.5583 | 15.475 | 15.7417 | 16.1 | 15.5167 | 15.5056 | 16.8833 | 17 | 14.6917 | 16.1367 | 14.5733 | 15.2 | 16.3333
| Median |  | 16.415 | 15.5149 | 17.443 | 16 | 16.1 | 15.55 | 15.45 | 15.85 | 16.025 | 15.375 | 15.4659 | 16.875 | 17 | 14.605 | 16.06 | 14.58 | 15.3 | 16.4
| Std Dev. |  | 0.21134 | 0.40772 | 0.028055 | 0.18348 | 0.23381 | 0.2417 | 0.30943 | 0.29397 | 0.54589 | 0.4446 | 0.47201 | 0.24728 | 0.18974 | 0.40642 | 0.2314 | 0.14514 | 0.36878 | 0.10328
| Rel.Std.Dev. |  | 1.291% | 2.616% | 0.1608% | 1.148% | 1.458% | 1.553% | 2% | 1.867% | 3.391% | 2.865% | 3.044% | 1.465% | 1.116% | 2.766% | 1.434% | 0.996% | 2.426% | 0.6323%
| PDM3 |  | 3.553% | -1.372% | 10.41% | 1.127% | 1.444% | -1.562% | -2.089% | -0.4016% | 1.866% | -1.825% | -1.895% | 6.822% | 7.56% | -7.045% | 2.098% | -7.794% | -3.829% | 3.342%
| Z-Score (Absolute) |  | 0.91 | 0.5 | 2.88 | 0.22 | 0.31 | 0.55 | 0.7 | 0.22 | 0.43 | 0.63 | 0.65 | 1.85 | 2.06 | 2.12 | 0.5 | 2.34 | 1.2 | 0.85
| Analytical results for Gd in OREAS 242b (Certified Value 3.35 ppm) | [21 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 3.47 | 3.4106 | 3.2 | 3.5 | 3.27 | 3.1 | 3.5 | 3.4 | [12 empty] | 1
| 1 | 2 | 3.47 | 3.3712 | 3.3 | 3.5 | 3.31 | 3.2 | 3.23 | 3.4 | [12 empty] | 6
| 1 | 3 | 3.52 | 3.5751 | 3.1 | 3.4 | 3.4 | 3.2 | 3.5 | 3.2 | [12 empty] | 16
| 1 | 4 | 3.55 | 3.41 | 3.2 | 3.6 | 3.29 | 3.2 | 3.38 | 3.2 | [12 empty] | 3.3479
| 1 | 5 | 3.32 | 3.398 | 3.1 | 3.4 | 3.3 | 3.1 | 3.42 | 3.4 | [12 empty] | 33
| 1 | 6 | 3.51 | 3.2827 | 3.1 | 3.4 | 3.43 | 3.2 | 3.4 | 3.4
| Mean |  | 3.4733 | 3.4079 | 3.1667 | 3.4667 | 3.3333 | 3.1667 | 3.405 | 3.3333
| Median |  | 3.49 | 3.404 | 3.15 | 3.45 | 3.305 | 3.2 | 3.41 | 3.4
| Std Dev. |  | 0.081158 | 0.094959 | 0.08165 | 0.08165 | 0.06532 | 0.05164 | 0.099549 | 0.10328
| Rel.Std.Dev. |  | 2.337% | 2.786% | 2.578% | 2.355% | 1.96% | 1.631% | 2.924% | 3.098%
| PDM3 |  | 3.745% | 1.792% | -5.415% | 3.546% | -0.4366% | -5.415% | 1.704% | -0.4366%
| Z-Score (Absolute) |  | 1.03 | 0.38 | 2 | 0.96 | 0.35 | 2 | 0.35 | 0.35
| Analytical results for Ge in OREAS 242b (Certified Value 0.13 ppm) | [21 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.6 | < 0.2 | 0.1565 | 0.3 | 0.07 | 0.14 | 0.11 | 0.09 | 0.19 | 0.99544 | NA | 1.4 | 0.14 | 0.9 | 1.32 | [5 empty] | 1
| 1 | 2 | 1.5 | < 0.2 | 0.1726 | 0.2 | 0.08 | 0.16 | 0.09 | 0.13 | 0.2 | 1.0158 | NA | 1.4 | 0.22 | 0.8 | 1.28 | [5 empty] | 38
| 1 | 3 | 1.5 | < 0.2 | 0.1474 | 0.2 | 0.08 | 0.16 | 0.1 | 0.09 | 0.13 | 1.2455 | NA | 1.3 | 0.16 | 1 | 1.32 | [5 empty] | 16
| 1 | 4 | 1.5 | < 0.2 | 0.1658 | 0.3 | 0.08 | 0.15 | 0.11 | 0.09 | 0.15 | 1.2709 | NA | 1.4 | 0.1 | 0.8 | 1.34 | [5 empty] | 0.12659
| 1 | 5 | 1.6 | < 0.2 | 0.1466 | 0.7 | 0.08 | 0.09 | 0.11 | 0.08 | 0.16 | 1.1442 | NA | 1.3 | 0.14 | 1 | 1.3 | [5 empty] | 34
| 1 | 6 | 1.5 | < 0.2 | 0.15 | 0.2 | 0.07 | 0.15 | 0.1 | 0.08 | 0.19 | 1.1213 | NA | 1.4 | 0.09 | 0.9 | 1.31
| Mean |  | 1.5333 |  | 0.15648 | 0.31667 | 0.076667 | 0.14167 | 0.10333 | 0.093333 | 0.17 | 1.1322 |  | 1.3667 | 0.14167 | 0.9 | 1.3117
| Median |  | 1.5 |  | 0.15325 | 0.25 | 0.08 | 0.15 | 0.105 | 0.09 | 0.175 | 1.1327 |  | 1.4 | 0.14 | 0.9 | 1.315
| Std Dev. |  | 0.05164 |  | 0.010666 | 0.19408 | 0.005164 | 0.026394 | 0.008165 | 0.018619 | 0.027568 | 0.11368 |  | 0.05164 | 0.046655 | 0.089443 | 0.020412
| Rel.Std.Dev. |  | 3.368% |  | 6.816% | 61.29% | 6.736% | 18.63% | 7.902% | 19.95% | 16.22% | 10.04% |  | 3.779% | 32.93% | 9.938% | 1.556%
| PDM3 |  | 1111% |  | 23.61% | 150.1% | -39.44% | 11.91% | -18.37% | -26.27% | 34.29% | 794.4% |  | 979.6% | 11.91% | 610.9% | 936.1%
| Z-Score (Absolute) |  | NA | 0.7 | 0.25 | NA | 1.1 | 0 | 0.65 | 0.81 | 0.48 | 16.7 | NA | NA | 0 | NA | 19.72
Results from laboratories 1.02, 1.07, 1.16 and 1.18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 242b (Certified Value 1.74 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 1.72 | 1.6769 | 1.1 | 2 | 1.8 | 1.8 | 1.8 | 1.8 | 1.7467 | 1.81 | 1.7 | 1.82 | 1.64 | 1.61 | 1.77 | 1.8 |  |  |  |  | 1
| 1 | 2 | 1.7 | 1.8398 | 0.6 | 2 | 1.8 | 1.7 | 1.9 | 1.9 | 1.7797 | 1.76 | 1.8 | 1.71 | 1.73 | 1.85 | 1.63 | 1.8 |  |  |  |  | 22
| 1 | 3 | 1.79 | 1.8047 | 0.6 | 1.8 | 1.8 | 1.8 | 2 | 1.8 | 1.7662 | 1.75 | 1.7 | 1.65 | 1.77 | 1.66 | 1.72 | 1.8 |  |  |  |  | 16
| 1 | 4 | 1.77 | 1.6645 | 1.5 | 1.7 | 1.8 | 1.8 | 1.8 | 1.8 | 1.7472 | 1.75 | 1.7 | 1.73 | 1.7 | 1.69 | 1.74 | 1.8 |  |  |  |  | 1.7365
| 1 | 5 | 1.59 | 1.6953 | 1.9 | 1.9 | 1.7 | 1.6 | 1.9 | 1.7 | 1.7713 | 1.76 | 1.7 | 1.78 | 1.71 | 1.57 | 1.8 | 1.8 |  |  |  |  | 35
| 1 | 6 | 1.65 | 1.6992 | 1.8 | 1.8 | 1.8 | 1.6 | 1.8 | 1.8 | 1.7586 | 1.78 | 1.7 | 1.68 | 1.69 | 1.62 | 1.72 | 1.8
| Mean |  | 1.7033 | 1.7301 | 1.25 | 1.8667 | 1.7833 | 1.7167 | 1.8667 | 1.8 | 1.7616 | 1.7683 | 1.7167 | 1.7283 | 1.7067 | 1.6667 | 1.73 | 1.8
| Median |  | 1.71 | 1.6973 | 1.3 | 1.85 | 1.8 | 1.75 | 1.85 | 1.8 | 1.7624 | 1.76 | 1.7 | 1.72 | 1.705 | 1.64 | 1.73 | 1.8
| Std Dev. |  | 0.074744 | 0.073353 | 0.57533 | 0.12111 | 0.040825 | 0.098319 | 0.08165 | 0.063246 | 0.013242 | 0.023166 | 0.040825 | 0.063061 | 0.043205 | 0.098928 | 0.057966 | 0
| Rel.Std.Dev. |  | 4.388% | 4.24% | 46.03% | 6.488% | 2.289% | 5.727% | 4.374% | 3.514% | 0.7517% | 1.31% | 2.378% | 3.649% | 2.532% | 5.936% | 3.351% | 0%
| PDM3 |  | -1.912% | -0.3727% | -28.02% | 7.493% | 2.695% | -1.144% | 7.493% | 3.654% | 1.444% | 1.831% | -1.144% | -0.4725% | -1.72% | -4.024% | -0.3766% | 3.654%
| Z-Score (Absolute) |  | 0.52 | 0 | 9.26 | 2.64 | 1.03 | 0.26 | 2.64 | 1.35 | 0.61 | 0.74 | 0.26 | 0.03 | 0.45 | 1.22 | 0 | 1.35
| Analytical results for Hg in OREAS 242b (Indicative Value 0.15 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.15 | 1.21 | [17 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [17 empty] | 2
2
| 1 | 1 | < 1 | 0.15 | 0.15 | [17 empty] | 1
| 1 | 2 | < 1 | 0.17 | 0.13 | [17 empty] | 2
| 1 | 3 | < 1 | 0.15 | 0.15 | [17 empty] | 16
| 1 | 4 | < 1 | 0.16 | 0.15 | [17 empty] | 0.15083
| 1 | 5 | < 1 | 0.15 | 0.15 | [17 empty] | 8
| 1 | 6 | < 1 | 0.16 | 0.14
| Mean |  |  | 0.15667 | 0.145
| Median |  |  | 0.155 | 0.15
| Std Dev. |  |  | 0.008165 | 0.0083666
| Rel.Std.Dev. |  |  | 5.212% | 5.77%
| PDM3 |  |  | 3.867% | -3.867%
| Z-Score (Absolute) |  | 19.84 | 0 | 0.67
| Analytical results for Ho in OREAS 242b (Certified Value 0.77 ppm) | [21 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.83 | 0.8162 | 0.92879 | 0.8 | 0.8 | 0.73 | 0.7 | 0.76 | 0.74 | [11 empty] | 1
| 1 | 2 | 0.8 | 0.77748 | 0.899 | 0.8 | 0.8 | 0.74 | 0.7 | 0.74 | 0.72 | [11 empty] | 23
| 1 | 3 | 0.81 | 0.80048 | 0.9188 | 0.8 | 0.8 | 0.75 | 0.7 | 0.78 | 0.72 | [11 empty] | 16
| 1 | 4 | 0.81 | 0.76714 | 0.95174 | 0.8 | 0.8 | 0.74 | 0.7 | 0.75 | 0.72 | [11 empty] | 0.76605
| 1 | 5 | 0.81 | 0.7648 | 0.93833 | 0.8 | 0.8 | 0.74 | 0.7 | 0.76 | 0.76 | [11 empty] | 36
| 1 | 6 | 0.84 | 0.75531 | 0.89702 | 0.8 | 0.7 | 0.73 | 0.7 | 0.78 | 0.74
| Mean |  | 0.81667 | 0.78023 | 0.92228 | 0.8 | 0.78333 | 0.73833 | 0.7 | 0.76167 | 0.73333
| Median |  | 0.81 | 0.77231 | 0.92379 | 0.8 | 0.8 | 0.74 | 0.7 | 0.76 | 0.73
| Std Dev. |  | 0.015055 | 0.023419 | 0.021726 | 1.2162e-16 | 0.040825 | 0.0075277 | 1.2162e-16 | 0.016021 | 0.01633
| Rel.Std.Dev. |  | 1.844% | 3.001% | 2.356% | 1.52e-14% | 5.212% | 1.02% | 1.737e-14% | 2.103% | 2.227%
| PDM3 |  | 6.608% | 1.852% | 20.39% | 4.432% | 2.257% | -3.618% | -8.622% | -0.5718% | -4.27%
| Z-Score (Absolute) |  | 0.88 | 0.18 | 2.91 | NA | NA | 0.63 | NA | 0.18 | 0.72
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 242b (Certified Value 0.074 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.07 | < 0.1 | < 2 | 0.082 | 0.078 | 0.08 | 0.073 | 0.077 | 0.073425 | 0.079 | 0.07 | 0.07 | 0.06 | 0.07 | 0.077 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.08 | < 0.1 | 3 | 0.083 | 0.083 | 0.074 | 0.082 | 0.08 | 0.071815 | 0.072 | 0.07 | 0.07 | 0.08 | 0.07 | 0.082 | 0.05 |  |  |  |  | 24
| 1 | 3 | 0.07 | < 0.1 | 4 | 0.072 | 0.077 | 0.079 | 0.075 | 0.083 | 0.067323 | 0.076 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.05 |  |  |  |  | 16
| 1 | 4 | 0.07 | < 0.1 | < 2 | 0.074 | 0.078 | 0.077 | 0.075 | 0.077 | 0.0771 | 0.074 | 0.07 | 0.07 | 0.08 | 0.07 | 0.08 | 0.05 |  |  |  |  | 0.07379
| 1 | 5 | 0.07 | < 0.1 | < 2 | 0.084 | 0.076 | 0.074 | 0.069 | 0.08 | 0.075775 | 0.067 | 0.07 | 0.07 | 0.07 | 0.07 | 0.064 | 0.05 |  |  |  |  | 37
| 1 | 6 | 0.07 | < 0.1 | < 2 | 0.078 | 0.081 | 0.073 | 0.071 | 0.079 | 0.068561 | 0.072 | 0.07 | 0.06 | 0.07 | 0.07 | 0.091 | 0.05
| Mean |  | 0.071667 |  | 3.5 | 0.078833 | 0.078833 | 0.076167 | 0.074167 | 0.079333 | 0.072333 | 0.073333 | 0.07 | 0.068333 | 0.071667 | 0.07 | 0.077333 | 0.05
| Median |  | 0.07 |  | 3.5 | 0.08 | 0.078 | 0.0755 | 0.074 | 0.0795 | 0.07262 | 0.073 | 0.07 | 0.07 | 0.07 | 0.07 | 0.0785 | 0.05
| Std Dev. |  | 0.0040825 |  | 0.70711 | 0.0049967 | 0.0026394 | 0.0029269 | 0.0044907 | 0.0022509 | 0.0038824 | 0.0040825 | 0 | 0.0040825 | 0.0075277 | 0 | 0.0094587 | 7.6012e-18
| Rel.Std.Dev. |  | 5.696% |  | 20.2% | 6.338% | 3.348% | 3.843% | 6.055% | 2.837% | 5.367% | 5.567% | 0% | 5.974% | 10.5% | 0% | 12.23% | 1.52e-14%
| PDM3 |  | -2.877% |  | 4643% | 6.835% | 6.835% | 3.221% | 0.5108% | 7.513% | -1.974% | -0.6185% | -5.136% | -7.395% | -2.877% | -5.136% | 4.802% | -32.24%
| Z-Score (Absolute) |  | 0.12 | 3.93 | NA | 1.14 | 1.14 | 0.67 | 0.32 | 1.23 | 0 | 0.18 | 0.41 | 0.7 | 0.12 | 0.41 | 0.88 | 3.93
| Analytical results for K in OREAS 242b (Certified Value 0.363 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.36 | 0.3793 | 0.3421 | 0.36 | 0.39 | 0.36 | 0.37 | 0.36 | 0.35 | 0.37 | 0.3661 | 0.4253 | 0.394 | 0.38 | 0.3391 | 0.34 | 0.362 | 0.35 | 0.35 | 0.3651 | 1
| 1 | 2 | 0.36 | 0.37484 | 0.34232 | 0.36 | 0.39 | 0.36 | 0.38 | 0.35 | 0.37 | 0.37 | 0.38087 | 0.4457 | 0.408 | 0.38 | 0.338 | 0.35 | 0.37 | 0.34 | 0.35 | 0.36311 | #N/A
| 1 | 3 | 0.3604 | 0.39076 | 0.34222 | 0.35 | 0.39 | 0.36 | 0.37 | 0.35 | 0.37 | 0.37 | 0.37737 | 0.4243 | 0.401 | 0.36 | 0.3476 | 0.34 | 0.371 | 0.34 | 0.35 | 1.627 | 16
| 1 | 4 | 0.3662 | 0.37942 | 0.34187 | 0.35 | 0.39 | 0.36 | 0.36 | 0.35 | 0.37 | 0.37 | 0.3854 | 0.4186 | 0.403 | 0.39 | 0.3367 | 0.35 | 0.37 | 0.34 | 0.35 | 0.37188 | 0.36313
| 1 | 5 | 0.3689 | 0.37106 | 0.34223 | 0.36 | 0.39 | 0.37 | 0.38 | 0.34 | 0.37 | 0.36 | 0.36522 | 0.4265 | 0.399 | 0.37 | 0.3412 | 0.35 | 0.335 | 0.34 | 0.36 | 0.36058 | 38
| 1 | 6 | 0.3649 | 0.38237 | 0.34259 | 0.36 | 0.38 | 0.37 | 0.37 | 0.35 | 0.37 | 0.36 | 0.37898 | 0.4248 | 0.396 | 0.36 | 0.3371 | 0.35 | 0.342 | 0.34 | 0.35 | 0.35655
| Mean |  | 0.3634 | 0.37963 | 0.34222 | 0.35667 | 0.38833 | 0.36333 | 0.37167 | 0.35 | 0.36667 | 0.36667 | 0.37566 | 0.42753 | 0.40017 | 0.37333 | 0.33995 | 0.34667 | 0.35833 | 0.34167 | 0.35167 | 0.57403
| Median |  | 0.36265 | 0.37936 | 0.34222 | 0.36 | 0.39 | 0.36 | 0.37 | 0.35 | 0.37 | 0.37 | 0.37818 | 0.42505 | 0.4 | 0.375 | 0.33855 | 0.35 | 0.366 | 0.34 | 0.35 | 0.3641
| Std Dev. |  | 0.0038068 | 0.0067529 | 0.00023919 | 0.005164 | 0.0040825 | 0.005164 | 0.0075277 | 0.0063246 | 0.008165 | 0.005164 | 0.008203 | 0.0093146 | 0.0050365 | 0.012111 | 0.0040815 | 0.005164 | 0.015858 | 0.0040825 | 0.0040825 | 0.51585
| Rel.Std.Dev. |  | 1.048% | 1.779% | 0.06989% | 1.448% | 1.051% | 1.421% | 2.025% | 1.807% | 2.227% | 1.408% | 2.184% | 2.179% | 1.259% | 3.244% | 1.201% | 1.49% | 4.425% | 1.195% | 1.161% | 89.86%
| PDM3 |  | 0.07442% | 4.543% | -5.758% | -1.78% | 6.941% | 0.05606% | 2.351% | -3.616% | 0.974% | 0.974% | 3.45% | 17.74% | 10.2% | 2.81% | -6.383% | -4.534% | -1.321% | -5.911% | -3.157% | 58.08%
| Z-Score (Absolute) |  | 0.08 | 0.7 | 1.1 | 0.4 | 1.12 | 0.08 | 0.32 | 0.73 | 0.08 | 0.08 | 0.51 | 3.01 | 1.69 | 0.4 | 1.21 | 0.89 | 0.32 | 1.13 | 0.64 | 10.08
| Analytical results for La in OREAS 242b (Certified Value 7.6 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  |  | 2
3
| 1 | 1 | 7.69 | 7.5926 | 8.7938 | 7.9 | 7.8 | 7.7 | 7.9 | 7.6 | 7.4 | 8.5 | 6.8 | 8 | 7.41 | 7.2 | 6.94 | 6.69 | 7.4 | 9.291 |  |  | 1
| 1 | 2 | 7.7 | 7.5744 | 8.7033 | 7.8 | 7.5 | 7.8 | 7.8 | 7.5 | 7.8 | 9 | 6.9 | 8.4 | 7.26 | 7.4 | 6.57 | 6.58 | 7.8 | 9.206 |  |  | 25
| 1 | 3 | 7.75 | 7.805 | 8.7692 | 7.9 | 7.4 | 7.5 | 7.6 | 7.5 | 8 | 8.2 | 6.8 | 7.7 | 7.59 | 7.4 | 6.94 | 6.74 | 7 | 44.031 |  |  | 16
| 1 | 4 | 7.79 | 7.6751 | 8.7247 | 8 | 7.6 | 7.4 | 7.7 | 7.4 | 7.9 | 8.5 | 6.8 | 8.2 | 7.31 | 7.2 | 7.01 | 6.78 | 7.3 | 9.084 |  |  | 7.6023
| 1 | 5 | 7.86 | 7.4091 | 8.7048 | 7.9 | 7.8 | 7.7 | 8.3 | 7.2 | 7.4 | 8.2 | 6.8 | 8.1 | 7.35 | 7.2 | 6.88 | 6.67 | 7.4 | 9.043 |  |  | 39
| 1 | 6 | 7.72 | 7.5488 | 8.7399 | 8.1 | 7.6 | 7.7 | 7.9 | 7.5 | 7.5 | 8.6 | 6.8 | 7.4 | 7.29 | 7.3 | 7.05 | 6.6 | 7.5 | 8.999
| Mean |  | 7.7517 | 7.6008 | 8.7393 | 7.9333 | 7.6167 | 7.6333 | 7.8667 | 7.45 | 7.6667 | 8.5 | 6.8167 | 7.9667 | 7.3683 | 7.2833 | 6.8983 | 6.6767 | 7.4 | 14.9423
| Median |  | 7.735 | 7.5835 | 8.7323 | 7.9 | 7.6 | 7.7 | 7.85 | 7.5 | 7.65 | 8.5 | 6.8 | 8.05 | 7.33 | 7.25 | 6.94 | 6.68 | 7.4 | 9.145
| Std Dev. |  | 0.064317 | 0.13224 | 0.036268 | 0.10328 | 0.16021 | 0.15055 | 0.24221 | 0.13784 | 0.26583 | 0.29665 | 0.040825 | 0.36148 | 0.1204 | 0.098319 | 0.17151 | 0.077632 | 0.26077 | 14.2509
| Rel.Std.Dev. |  | 0.8297% | 1.74% | 0.415% | 1.302% | 2.103% | 1.972% | 3.079% | 1.85% | 3.467% | 3.49% | 0.5989% | 4.537% | 1.634% | 1.35% | 2.486% | 1.163% | 3.524% | 95.37%
| PDM3 |  | 1.964% | -0.01995% | 14.96% | 4.354% | 0.1885% | 0.4077% | 3.477% | -2.004% | 0.8462% | 11.81% | -10.33% | 4.792% | -3.078% | -4.196% | -9.26% | -12.18% | -2.662% | 96.55%
| Z-Score (Absolute) |  | 0.3 | 0.06 | 2.66 | 0.74 | 0.02 | 0.02 | 0.58 | 0.42 | 0.1 | 2.09 | 1.93 | 0.82 | 0.61 | 0.82 | 1.73 | 2.26 | 0.54 | 17.47
| Analytical results for Li in OREAS 242b (Certified Value 19.7 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 19.1 | 19.3068 | 9.8766 | 20 | 17.9 | 21.2 | 20.4 | 21.6 | 20.1 | 21.3 | 19.0788 | 15 | 20.1 | 18.1 | 18.6 | 16.6 | 19.5 | 21.81 | 20 | 19.462 | 1
| 1 | 2 | 19.3 | 19.4111 | 9.7858 | 20 | 18.2 | 21.1 | 20.9 | 21.3 | 20.8 | 21.6 | 19.5981 | 15 | 20.1 | 17.2 | 18.4 | 17.2 | 18.8 | 21.91 | 20 | 19.331 | #N/A
| 1 | 3 | 18.9 | 20.111 | 9.3131 | 20 | 17.9 | 20.8 | 21 | 21.5 | 21.1 | 21.5 | 19.0619 | 15 | 19.1 | 17.4 | 19 | 16.9 | 20.5 | 21.16 | 20 | 29.049 | 16
| 1 | 4 | 19 | 19.3661 | 9.8116 | 20 | 19.2 | 20.9 | 20.5 | 20.8 | 21.4 | 21.3 | 19.2313 | 14 | 20.1 | 17.3 | 18.6 | 17.2 | 20.5 | 21.09 | 20 | 19.499 | 19.6618
| 1 | 5 | 19.2 | 19.3557 | 9.7955 | 20 | 18.5 | 21 | 20.6 | 20.5 | 21.1 | 21 | 19.5977 | 15 | 20.1 | 16.4 | 18.4 | 16.9 | 19.9 | 19.94 | 20 | 19.222 | 40
| 1 | 6 | 18.9 | 19.6261 | 9.7833 | 20 | 17.8 | 21.2 | 21 | 21 | 21.2 | 20.9 | 19.5712 | 15 | 20.1 | 16.9 | 18.5 | 17.2 | 20.1 | 20.44 | 20 | 18.853
| Mean |  | 19.0667 | 19.5295 | 9.7277 | 20 | 18.25 | 21.0333 | 20.7333 | 21.1167 | 20.95 | 21.2667 | 19.3565 | 14.8333 | 19.9333 | 17.2167 | 18.5833 | 17 | 19.8833 | 21.0583 | 20 | 20.9027
| Median |  | 19.05 | 19.3886 | 9.7906 | 20 | 18.05 | 21.05 | 20.75 | 21.15 | 21.1 | 21.3 | 19.4013 | 15 | 20.1 | 17.25 | 18.55 | 17.05 | 20 | 21.125 | 20 | 19.3965
| Std Dev. |  | 0.1633 | 0.30593 | 0.206 | 0 | 0.53198 | 0.1633 | 0.26583 | 0.42622 | 0.45935 | 0.27325 | 0.26162 | 0.40825 | 0.40825 | 0.56362 | 0.22286 | 0.24495 | 0.65243 | 0.76612 | 0 | 3.9976
| Rel.Std.Dev. |  | 0.8565% | 1.566% | 2.118% | 0% | 2.915% | 0.7764% | 1.282% | 2.018% | 2.193% | 1.285% | 1.352% | 2.752% | 2.048% | 3.274% | 1.199% | 1.441% | 3.281% | 3.638% | 0% | 19.12%
| PDM3 |  | -3.027% | -0.6732% | -50.53% | 1.72% | -7.181% | 6.975% | 5.45% | 7.399% | 6.552% | 8.162% | -1.553% | -24.56% | 1.381% | -12.44% | -5.485% | -13.54% | 1.127% | 7.103% | 1.72% | 6.311%
| Z-Score (Absolute) |  | 0.52 | 0.24 | 6.34 | 0.06 | 1.03 | 0.7 | 0.51 | 0.75 | 0.65 | 0.85 | 0.34 | NA | 0.02 | 1.68 | 0.82 | 1.81 | 0.02 | 0.72 | NA | 0.62
Results from laboratory 1.15 were removed due to their 1 ppm reading resolution. Results from laboratory 1.28 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 242b (Certified Value 0.32 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.34 | 0.33372 | 0.39482 | 0.3 | 0.3 | 0.32 | 0.3 | 0.33 | 0.32 | [11 empty] | 1
| 1 | 2 | 0.35 | 0.33993 | 0.35657 | 0.3 | 0.3 | 0.31 | 0.3 | 0.31 | 0.32 | [11 empty] | 7
| 1 | 3 | 0.32 | 0.32094 | 0.39032 | 0.3 | 0.3 | 0.31 | 0.3 | 0.32 | 0.34 | [11 empty] | 16
| 1 | 4 | 0.33 | 0.31292 | 0.38546 | 0.3 | 0.3 | 0.32 | 0.3 | 0.31 | 0.34 | [11 empty] | 0.32405
| 1 | 5 | 0.34 | 0.32161 | 0.38789 | 0.4 | 0.3 | 0.32 | 0.3 | 0.3 | 0.32 | [11 empty] | 41
| 1 | 6 | 0.38 | 0.29642 | 0.36746 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.34
| Mean |  | 0.34333 | 0.32092 | 0.38042 | 0.31667 | 0.3 | 0.32 | 0.3 | 0.31333 | 0.33
| Median |  | 0.34 | 0.32128 | 0.38667 | 0.3 | 0.3 | 0.32 | 0.3 | 0.31 | 0.33
| Std Dev. |  | 0.020656 | 0.015425 | 0.014989 | 0.040825 | 0 | 0.010954 | 0 | 0.010328 | 0.010954
| Rel.Std.Dev. |  | 6.016% | 4.806% | 3.94% | 12.89% | 0% | 3.423% | 0% | 3.296% | 3.32%
| PDM3 |  | 5.95% | -0.965% | 17.39% | -2.279% | -7.422% | -1.25% | -7.422% | -3.308% | 1.836%
| Z-Score (Absolute) |  | 1.37 | 0.35 | 4.21 | NA | NA | 0.42 | NA | 0.93 | 0.35
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 242b (Certified Value 3.52 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.5509 | 3.4214 | 3.5421 | 3.5 | 3.45 | 3.45 | 3.49 | 3.56 | 3.26 | 3.31 | 3.5907 | 3.5756 | 3.52 | 3.48 | 3.1503 | 3.38 | 3.78 | 3.56 | 3.65 | 3.6704 | 1
| 1 | 2 | 3.4836 | 3.3653 | 3.5457 | 3.47 | 3.39 | 3.44 | 3.51 | 3.5 | 3.51 | 3.33 | 3.6523 | 3.65 | 3.54 | 3.58 | 3.1491 | 3.47 | 3.54 | 3.46 | 3.61 | 3.6496 | #N/A
| 1 | 3 | 3.531 | 3.517 | 3.5384 | 3.43 | 3.4 | 3.39 | 3.49 | 3.46 | 3.42 | 3.33 | 3.6197 | 3.6059 | 3.48 | 3.37 | 3.2257 | 3.42 | 3.34 | 3.53 | 3.59 | 1.4523 | 16
| 1 | 4 | 3.5363 | 3.4176 | 3.5359 | 3.49 | 3.43 | 3.46 | 3.47 | 3.43 | 3.43 | 3.3 | 3.6791 | 3.5791 | 3.6 | 3.63 | 3.1672 | 3.47 | 3.59 | 3.52 | 3.62 | 3.676 | 3.5158
| 1 | 5 | 3.5944 | 3.3569 | 3.531 | 3.49 | 3.41 | 3.45 | 3.49 | 3.39 | 3.43 | 3.27 | 3.5849 | 3.5641 | 3.57 | 3.52 | 3.1878 | 3.46 | 3.21 | 3.5 | 3.7 | 3.6231 | 42
| 1 | 6 | 3.538 | 3.4372 | 3.5385 | 3.52 | 3.43 | 3.45 | 3.51 | 3.49 | 3.44 | 3.24 | 3.6647 | 3.6338 | 3.48 | 3.35 | 3.1516 | 3.46 | 3.46 | 3.5 | 3.62 | 3.6457
| Mean |  | 3.539 | 3.4192 | 3.5386 | 3.4833 | 3.4183 | 3.44 | 3.4933 | 3.4717 | 3.415 | 3.2967 | 3.6319 | 3.6014 | 3.5317 | 3.4883 | 3.172 | 3.4433 | 3.4867 | 3.5117 | 3.6317 | 3.2862
| Median |  | 3.5372 | 3.4195 | 3.5385 | 3.49 | 3.42 | 3.45 | 3.49 | 3.475 | 3.43 | 3.305 | 3.636 | 3.5925 | 3.53 | 3.5 | 3.1594 | 3.46 | 3.5 | 3.51 | 3.62 | 3.6477
| Std Dev. |  | 0.035643 | 0.057766 | 0.0050759 | 0.030768 | 0.022286 | 0.025298 | 0.015055 | 0.059133 | 0.082644 | 0.03559 | 0.039421 | 0.034601 | 0.048339 | 0.11197 | 0.030225 | 0.036148 | 0.19916 | 0.033714 | 0.038687 | 0.89861
| Rel.Std.Dev. |  | 1.007% | 1.689% | 0.1434% | 0.8833% | 0.652% | 0.7354% | 0.431% | 1.703% | 2.42% | 1.08% | 1.085% | 0.9608% | 1.369% | 3.21% | 0.9529% | 1.05% | 5.712% | 0.9601% | 1.065% | 27.34%
| PDM3 |  | 0.6607% | -2.747% | 0.6486% | -0.9236% | -2.772% | -2.156% | -0.6392% | -1.255% | -2.867% | -6.233% | 3.302% | 2.435% | 0.4512% | -0.7814% | -9.78% | -2.061% | -0.8288% | -0.1177% | 3.295% | -6.531%
| Z-Score (Absolute) |  | 0.68 | 0.82 | 0.67 | 0.02 | 0.84 | 0.56 | 0.1 | 0.17 | 0.88 | 2.36 | 1.84 | 1.46 | 0.58 | 0.04 | 3.92 | 0.52 | 0.02 | 0.33 | 1.84 | 2.49
| Analytical results for Mn in OREAS 242b (Certified Value 0.126 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.1312 | 0.1253 | 0.13334 | 0.127 | 0.142 | 0.1225 | 0.121 | 0.131 | 0.118 | 0.1215 | 0.12328 | 0.1115 | 0.127 | 0.1246 | 0.1205 | 0.1235 | 0.14 | 0.1278 | 0.122 | 0.12706 | 1
| 1 | 2 | 0.1289 | 0.12288 | 0.13309 | 0.123 | 0.1334 | 0.121 | 0.1225 | 0.1265 | 0.129 | 0.122 | 0.12931 | 0.1134 | 0.132 | 0.1274 | 0.1218 | 0.128 | 0.131 | 0.1251 | 0.117 | 0.12444 | 27
| 1 | 3 | 0.1303 | 0.13008 | 0.13287 | 0.123 | 0.1346 | 0.1205 | 0.121 | 0.126 | 0.1235 | 0.122 | 0.1266 | 0.112 | 0.129 | 0.1208 | 0.1248 | 0.1252 | 0.123 | 0.1271 | 0.12 | 0.038026 | 16
| 1 | 4 | 0.1307 | 0.12477 | 0.1331 | 0.127 | 0.1372 | 0.122 | 0.12 | 0.1245 | 0.1225 | 0.1205 | 0.12912 | 0.1111 | 0.13 | 0.1299 | 0.1218 | 0.127 | 0.133 | 0.1267 | 0.119 | 0.12992 | 0.12605
| 1 | 5 | 0.1329 | 0.12288 | 0.13247 | 0.127 | 0.1373 | 0.1225 | 0.1265 | 0.1215 | 0.1235 | 0.1205 | 0.12672 | 0.1111 | 0.129 | 0.1254 | 0.1222 | 0.1262 | 0.118 | 0.1258 | 0.123 | 0.12353 | 43
| 1 | 6 | 0.131 | 0.12708 | 0.13367 | 0.127 | 0.1363 | 0.1225 | 0.1225 | 0.1275 | 0.1255 | 0.119 | 0.12231 | 0.1127 | 0.129 | 0.12 | 0.1206 | 0.127 | 0.128 | 0.1257 | 0.12 | 0.12316
| Mean |  | 0.13083 | 0.1255 | 0.13309 | 0.12567 | 0.1368 | 0.12183 | 0.12225 | 0.12617 | 0.12367 | 0.12092 | 0.12622 | 0.11197 | 0.12933 | 0.12468 | 0.12195 | 0.12615 | 0.12883 | 0.12637 | 0.12017 | 0.11102
| Median |  | 0.13085 | 0.12504 | 0.13309 | 0.127 | 0.13675 | 0.12225 | 0.12175 | 0.12625 | 0.1235 | 0.121 | 0.12666 | 0.11175 | 0.129 | 0.125 | 0.1218 | 0.1266 | 0.1295 | 0.12625 | 0.12 | 0.12398
| Std Dev. |  | 0.0013018 | 0.0027501 | 0.00040789 | 0.0020656 | 0.0029698 | 0.0008756 | 0.0022967 | 0.003157 | 0.0036148 | 0.0011583 | 0.0029076 | 0.0009288 | 0.001633 | 0.0037981 | 0.0015592 | 0.0015997 | 0.0077309 | 0.0010073 | 0.002137 | 0.035851
| Rel.Std.Dev. |  | 0.995% | 2.191% | 0.3065% | 1.644% | 2.171% | 0.7187% | 1.879% | 2.502% | 2.923% | 0.9579% | 2.304% | 0.8295% | 1.263% | 3.046% | 1.279% | 1.268% | 6.001% | 0.7971% | 1.778% | 32.29%
| PDM3 |  | 3.792% | -0.4392% | 5.582% | -0.3066% | 8.526% | -3.348% | -3.017% | 0.09011% | -1.893% | -4.075% | 0.1354% | -11.17% | 2.602% | -1.087% | -3.255% | 0.07688% | 2.206% | 0.2488% | -4.67% | -11.93%
| Z-Score (Absolute) |  | 1.02 | 0.02 | 1.45 | 0.02 | 2.17 | 0.73 | 0.65 | 0.11 | 0.37 | 0.9 | 0.12 | 2.64 | 0.73 | 0.17 | 0.7 | 0.11 | 0.63 | 0.15 | 1.05 | 2.82
| Analytical results for Mo in OREAS 242b (Certified Value 8.2 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 8.6 | 7.5716 | 2 | 8.8 | 8.44 | 8.43 | 9.3 | 8.32 | 8.36 | 7.6973 | 8.1 | 8.5 | 7.88 | 8.4 | 7.52 | 7.9 | 6.72 | 8.5 | 8.245 |  | 1
| 1 | 2 | 8.5 | 7.4113 | 1 | 8 | 8.44 | 8.36 | 9.03 | 8.66 | 8.48 | 7.8171 | 8 | 8.3 | 7.45 | 8.3 | 7.67 | 7.6 | 6.01 | 8.5 | 8.52 |  | 28
| 1 | 3 | 8.6 | 7.7255 | < 1 | 8.3 | 8.23 | 8.52 | 9.17 | 9.02 | 8.55 | 7.7139 | 7.8 | 8.4 | 7.66 | 8.9 | 7.57 | 7.6 | 5.59 | 8.5 | 2.97 |  | 16
| 1 | 4 | 8.8 | 7.4618 | 5 | 8.1 | 7.91 | 8.36 | 9.02 | 8.6 | 8.59 | 7.5819 | 7.8 | 8.4 | 7.55 | 8.5 | 7.69 | 8.1 | 6.86 | 8.5 | 8.267 |  | 8.1966
| 1 | 5 | 8.7 | 7.2098 | 8 | 8.3 | 8.54 | 8.31 | 8.76 | 8.33 | 8.14 | 7.8321 | 7.8 | 8.4 | 7.41 | 8.5 | 7.52 | 8.1 | 6.66 | 8.5 | 8.301 |  | 44
| 1 | 6 | 8.8 | 7.5664 | 8 | 8.3 | 8.16 | 8.62 | 9.02 | 8.28 | 8.44 | 7.5747 | 8 | 8.2 | 7.18 | 8.1 | 7.72 | 8.1 | 6.96 | 8.5 | 7.565
| Mean |  | 8.6667 | 7.4911 | 4.8 | 8.3 | 8.2867 | 8.4333 | 9.05 | 8.535 | 8.4267 | 7.7028 | 7.9167 | 8.3667 | 7.5217 | 8.45 | 7.615 | 7.9 | 6.4667 | 8.5 | 7.3113
| Median |  | 8.65 | 7.5141 | 5 | 8.3 | 8.335 | 8.395 | 9.025 | 8.465 | 8.46 | 7.7056 | 7.9 | 8.4 | 7.5 | 8.45 | 7.62 | 8 | 6.69 | 8.5 | 8.256
| Std Dev. |  | 0.12111 | 0.17521 | 3.2711 | 0.27568 | 0.23338 | 0.1169 | 0.18067 | 0.28578 | 0.16219 | 0.11041 | 0.13292 | 0.10328 | 0.23777 | 0.26646 | 0.089163 | 0.24495 | 0.54346 | 0 | 2.1511
| Rel.Std.Dev. |  | 1.397% | 2.339% | 68.15% | 3.321% | 2.816% | 1.386% | 1.996% | 3.348% | 1.925% | 1.433% | 1.679% | 1.234% | 3.161% | 3.153% | 1.171% | 3.101% | 8.404% | 0% | 29.42%
| PDM3 |  | 5.735% | -8.607% | -41.44% | 1.262% | 1.099% | 2.888% | 10.41% | 4.129% | 2.807% | -6.024% | -3.415% | 2.075% | -8.234% | 3.092% | -7.095% | -3.618% | -21.11% | 3.702% | -10.8%
| Z-Score (Absolute) |  | 0.67 | 1.44 | NA | 0.01 | 0.01 | 0.25 | 1.36 | 0.43 | 0.24 | 1.06 | 0.68 | 0.13 | 1.39 | 0.28 | 1.22 | 0.71 | 3.28 | 0.37 | 1.77
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 242b (Certified Value 1.73 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.7853 | 1.7439 | 1.5929 | 1.67 | 1.78 | 1.72 | 1.71 | 1.76 | 1.6 | 1.75 | 1.7564 | 1.8054 | 1.81 | 1.759 | 1.5993 | 1.646 | 1.85 | 1.74 | 1.67 | 1.7438 | 1
| 1 | 2 | 1.764 | 1.6988 | 1.5826 | 1.61 | 1.75 | 1.69 | 1.74 | 1.74 | 1.76 | 1.76 | 1.7549 | 1.8366 | 1.84 | 1.742 | 1.6102 | 1.714 | 1.74 | 1.74 | 1.68 | 1.737 | #N/A
| 1 | 3 | 1.7704 | 1.7773 | 1.5654 | 1.66 | 1.75 | 1.68 | 1.73 | 1.72 | 1.7 | 1.76 | 1.746 | 1.8335 | 1.79 | 1.657 | 1.6482 | 1.686 | 1.63 | 1.72 | 1.68 | 0.97495 | 16
| 1 | 4 | 1.7832 | 1.7312 | 1.5784 | 1.71 | 1.81 | 1.69 | 1.69 | 1.71 | 1.71 | 1.74 | 1.7318 | 1.8005 | 1.84 | 1.777 | 1.6065 | 1.695 | 1.75 | 1.8 | 1.67 | 1.7789 | 1.7307
| 1 | 5 | 1.8131 | 1.7236 | 1.5654 | 1.67 | 1.76 | 1.68 | 1.75 | 1.69 | 1.68 | 1.71 | 1.7347 | 1.79 | 1.85 | 1.717 | 1.6184 | 1.682 | 1.78 | 1.76 | 1.71 | 1.7245 | 45
| 1 | 6 | 1.7988 | 1.7453 | 1.5656 | 1.55 | 1.78 | 1.7 | 1.74 | 1.72 | 1.7 | 1.72 | 1.7567 | 1.8302 | 1.81 | 1.684 | 1.603 | 1.698 | 1.7 | 1.79 | 1.68 | 1.7136
| Mean |  | 1.7858 | 1.7367 | 1.5751 | 1.645 | 1.7717 | 1.6933 | 1.7267 | 1.7233 | 1.6917 | 1.74 | 1.7468 | 1.816 | 1.8233 | 1.7227 | 1.6143 | 1.6868 | 1.7417 | 1.7583 | 1.6817 | 1.6121
| Median |  | 1.7843 | 1.7375 | 1.572 | 1.665 | 1.77 | 1.69 | 1.735 | 1.72 | 1.7 | 1.745 | 1.7505 | 1.8178 | 1.825 | 1.7295 | 1.6084 | 1.6905 | 1.745 | 1.75 | 1.68 | 1.7308
| Std Dev. |  | 0.018076 | 0.026141 | 0.011509 | 0.05648 | 0.023166 | 0.015055 | 0.022509 | 0.024221 | 0.052313 | 0.020976 | 0.011197 | 0.019803 | 0.023381 | 0.045794 | 0.017868 | 0.022895 | 0.07414 | 0.031252 | 0.01472 | 0.31295
| Rel.Std.Dev. |  | 1.012% | 1.505% | 0.7307% | 3.433% | 1.308% | 0.8891% | 1.304% | 1.405% | 3.092% | 1.206% | 0.641% | 1.09% | 1.282% | 2.658% | 1.107% | 1.357% | 4.257% | 1.777% | 0.8753% | 19.41%
| PDM3 |  | 3.186% | 0.3476% | -8.991% | -4.95% | 2.369% | -2.157% | -0.231% | -0.4236% | -2.253% | 0.5395% | 0.9296% | 4.933% | 5.355% | -0.4621% | -6.726% | -2.533% | 0.6358% | 1.599% | -2.831% | -6.848%
| Z-Score (Absolute) |  | 1.15 | 0.22 | 2.83 | 1.51 | 0.88 | 0.6 | 0.03 | 0.03 | 0.63 | 0.28 | 0.41 | 1.72 | 1.85 | 0.04 | 2.09 | 0.72 | 0.31 | 0.63 | 0.82 | 2.13
| Analytical results for Nb in OREAS 242b (Certified Value 3.89 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 3.96 | 3.8635 | 3.6 | < 5 | 4.1 | 4.1 | 4.1 | 3.9 | 3.9 | 4.0392 | NA | 4.1 | 3.85 | 3.94 | 3.51 | 3.9 | 3.5 |  |  |  | 1
| 1 | 2 | 3.94 | 3.7833 | 3.6 | < 5 | 4 | 4.1 | 4 | 4.1 | 3.9 | 4.0796 | NA | 4.1 | 3.65 | 3.94 | 3.6 | 3.8 | 3.5 |  |  |  | 29
| 1 | 3 | 4.07 | 3.9031 | 3.6 | < 5 | 3.9 | 4.1 | 4 | 4.2 | 3.9 | 4.0204 | NA | 4.1 | 3.67 | 3.89 | 3.57 | 4 | 4 |  |  |  | 16
| 1 | 4 | 4.08 | 3.7133 | 3.8 | < 5 | 3.8 | 4 | 3.9 | 4.3 | 3.8 | 4.057 | NA | 4 | 3.83 | 3.92 | 3.56 | 3.9 | 3.5 |  |  |  | 3.8943
| 1 | 5 | 3.99 | 3.7444 | 3.7 | < 5 | 3.9 | 3.8 | 3.8 | 4.1 | 3.8 | 4.0825 | NA | 4 | 3.73 | 3.96 | 3.55 | 3.9 | 4 |  |  |  | 46
| 1 | 6 | 4.15 | 3.8686 | 3.7 | < 5 | 4 | 4.2 | 4 | 4 | 3.9 | 4.0439 | NA | 4.1 | 3.7 | 3.93 | 3.6 | 3.9 | 3.5
| Mean |  | 4.0317 | 3.8127 | 3.6667 |  | 3.95 | 4.05 | 3.9667 | 4.1 | 3.8667 | 4.0538 |  | 4.0667 | 3.7383 | 3.93 | 3.565 | 3.9 | 3.6667
| Median |  | 4.03 | 3.8234 | 3.65 |  | 3.95 | 4.1 | 4 | 4.1 | 3.9 | 4.0504 |  | 4.1 | 3.715 | 3.935 | 3.565 | 3.9 | 3.5
| Std Dev. |  | 0.081343 | 0.076553 | 0.08165 |  | 0.10488 | 0.13784 | 0.10328 | 0.14142 | 0.05164 | 0.024177 |  | 0.05164 | 0.083526 | 0.023664 | 0.033912 | 0.063246 | 0.2582
| Rel.Std.Dev. |  | 2.018% | 2.008% | 2.227% |  | 2.655% | 3.403% | 2.604% | 3.449% | 1.336% | 0.5964% |  | 1.27% | 2.234% | 0.6021% | 0.9512% | 1.622% | 7.042%
| PDM3 |  | 3.527% | -2.095% | -5.846% |  | 1.43% | 3.998% | 1.858% | 5.282% | -0.7101% | 4.094% |  | 4.426% | -4.005% | 0.9162% | -8.456% | 0.1459% | -5.846%
| Z-Score (Absolute) |  | 0.57 | 0.5 | 1.22 | 6.97 | 0.17 | 0.67 | 0.25 | 0.91 | 0.24 | 0.68 | NA | 0.75 | 0.87 | 0.07 | 1.72 | 0.07 | 1.22
| Analytical results for Nd in OREAS 242b (Certified Value 9.9 ppm) | [21 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 10.09 | 10.333 | 12.2621 | 10.5 | 10.3 | 9.79 | 9.3 | 9.8 | 10.4 | [11 empty] | 1
| 1 | 2 | 10.03 | 9.7529 | 12.3089 | 10.2 | 9.7 | 9.53 | 9.6 | 9.61 | 10.4 | [11 empty] | 4
| 1 | 3 | 10 | 10.1714 | 12.256 | 10.1 | 10.1 | 9.67 | 9.4 | 9.93 | 10.3 | [11 empty] | 16
| 1 | 4 | 10.1 | 9.8025 | 12.2811 | 10.3 | 9.9 | 9.64 | 9.3 | 9.8 | 10 | [11 empty] | 9.9039
| 1 | 5 | 10.02 | 9.9656 | 12.282 | 10.4 | 9.8 | 9.62 | 9.2 | 9.73 | 10.2 | [11 empty] | 47
| 1 | 6 | 10 | 9.8307 | 12.312 | 10.4 | 9.6 | 9.72 | 9.3 | 10.1 | 9.95
| Mean |  | 10.04 | 9.976 | 12.2837 | 10.3167 | 9.9 | 9.6617 | 9.35 | 9.8283 | 10.2083
| Median |  | 10.025 | 9.8981 | 12.2815 | 10.35 | 9.85 | 9.655 | 9.3 | 9.8 | 10.25
| Std Dev. |  | 0.044272 | 0.23098 | 0.023185 | 0.1472 | 0.26077 | 0.088863 | 0.13784 | 0.16916 | 0.196
| Rel.Std.Dev. |  | 0.441% | 2.315% | 0.1887% | 1.427% | 2.634% | 0.9198% | 1.474% | 1.721% | 1.92%
| PDM3 |  | 1.374% | 0.7283% | 24.03% | 4.168% | -0.03913% | -2.446% | -5.593% | -0.7628% | 3.074%
| Z-Score (Absolute) |  | 0.19 | 0 | 6.7 | 0.99 | 0.22 | 0.91 | 1.82 | 0.43 | 0.67
| Analytical results for Ni in OREAS 242b (Certified Value 75 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 77.6 | 76.1025 | 81.1578 | 74 | 75 | 80.2 | 77.2 | 77.7 | 76.5 | 75.2 | 72.9958 | 69.1 | 75 | 70.7 | 73.1 | 70.3 | 75.8 | 71.94 | 88 | 74.131 | 1
| 1 | 2 | 76.2 | 74.4985 | 80.788 | 74 | 74 | 78 | 77.8 | 77.8 | 76.3 | 78.8 | 76.4988 | 67.9 | 75 | 72.4 | 72.3 | 72.5 | 75 | 72.71 | 88 | 73.673 | 14
| 1 | 3 | 76.9 | 77.2844 | 80.4298 | 71 | 74 | 75.9 | 77.3 | 77 | 81.7 | 74.6 | 74.5135 | 67.5 | 75 | 71.4 | 74.3 | 71.9 | 75 | 69.93 | 88 | 72.383 | 16
| 1 | 4 | 77.9 | 75.6816 | 79.0662 | 74 | 76 | 75.1 | 75.5 | 76.5 | 83.4 | 75.3 | 77.3265 | 66.8 | 76 | 73.7 | 73.2 | 72 | 78.4 | 70.39 | 88 | 74.993 | 75.2484
| 1 | 5 | 77.8 | 74.0091 | 79.4475 | 74 | 75 | 76.9 | 75.4 | 74.1 | 79.5 | 74 | 75.6609 | 66.5 | 76 | 71.3 | 73.5 | 72.1 | 74 | 70.03 | 90 | 74.504 | 48
| 1 | 6 | 77.5 | 75.8264 | 79.2279 | 73 | 75 | 79.3 | 79.3 | 77.3 | 78.8 | 75.2 | 72.6962 | 66.1 | 75 | 68.1 | 72.4 | 72.3 | 76.1 | 69.46 | 88 | 73.157
| Mean |  | 77.3167 | 75.5671 | 80.0195 | 73.3333 | 74.8333 | 77.5667 | 77.0833 | 76.7333 | 79.3667 | 75.5167 | 74.9486 | 67.3167 | 75.3333 | 71.2667 | 73.1333 | 71.85 | 75.7167 | 70.7433 | 88.3333 | 73.8068
| Median |  | 77.55 | 75.754 | 79.9386 | 74 | 75 | 77.45 | 77.25 | 77.15 | 79.15 | 75.2 | 75.0872 | 67.15 | 75 | 71.35 | 73.15 | 72.05 | 75.4 | 70.21 | 88 | 73.902
| Std Dev. |  | 0.64936 | 1.1735 | 0.88512 | 1.2111 | 0.75277 | 1.9715 | 1.4716 | 1.375 | 2.8169 | 1.6833 | 1.8784 | 1.0926 | 0.5164 | 1.8747 | 0.73937 | 0.7893 | 1.5052 | 1.2839 | 0.8165 | 0.94505
| Rel.Std.Dev. |  | 0.8399% | 1.553% | 1.106% | 1.651% | 1.006% | 2.542% | 1.909% | 1.792% | 3.549% | 2.229% | 2.506% | 1.623% | 0.6855% | 2.631% | 1.011% | 1.099% | 1.988% | 1.815% | 0.9243% | 1.28%
| PDM3 |  | 2.749% | 0.4234% | 6.34% | -2.545% | -0.5517% | 3.081% | 2.438% | 1.973% | 5.473% | 0.3564% | -0.3985% | -10.54% | 0.1128% | -5.292% | -2.811% | -4.516% | 0.6222% | -5.987% | 17.39% | -1.916%
| Z-Score (Absolute) |  | 0.64 | 0.05 | 1.56 | 0.71 | 0.2 | 0.73 | 0.56 | 0.44 | 1.33 | 0.03 | 0.16 | 2.75 | 0.03 | 1.41 | 0.78 | 1.21 | 0.1 | 1.59 | 4.37 | 0.55
| Analytical results for P in OREAS 242b (Certified Value 0.04 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.0407 | 0.040047 | 0.035 | 0.0415 | 0.041 | 0.039 | 0.039 | 0.039 | 0.04 | 0.040916 | 0.0403 | 0.041 | 0.04 | 0.0394 | 0.036 | 0.0262 | 0.04 | 0.04 | 0.040441 |  | 1
| 1 | 2 | 0.0398 | 0.039329 | 0.034 | 0.0421 | 0.041 | 0.04 | 0.039 | 0.043 | 0.04 | 0.042358 | 0.0409 | 0.042 | 0.041 | 0.0391 | 0.037 | 0.0257 | 0.04 | 0.04 | 0.0406 |  | #N/A
| 1 | 3 | 0.0403 | 0.040782 | 0.035 | 0.0407 | 0.04 | 0.04 | 0.038 | 0.041 | 0.041 | 0.040298 | 0.0405 | 0.041 | 0.039 | 0.0376 | 0.037 | 0.0278 | 0.05 | 0.04 | 0.094625 |  | 16
| 1 | 4 | 0.0406 | 0.04027 | 0.038 | 0.0414 | 0.041 | 0.039 | 0.039 | 0.041 | 0.04 | 0.042893 | 0.0387 | 0.042 | 0.042 | 0.0364 | 0.038 | 0.0331 | 0.04 | 0.04 | 0.0408 |  | 0.039944
| 1 | 5 | 0.0411 | 0.038695 | 0.042 | 0.0417 | 0.041 | 0.04 | 0.038 | 0.042 | 0.04 | 0.039541 | 0.0398 | 0.042 | 0.04 | 0.04 | 0.036 | 0.028 | 0.05 | 0.04 | 0.039881 |  | 49
| 1 | 6 | 0.0407 | 0.039892 | 0.04 | 0.0407 | 0.042 | 0.04 | 0.04 | 0.042 | 0.04 | 0.042442 | 0.0409 | 0.041 | 0.039 | 0.0393 | 0.037 | 0.0317 | 0.04 | 0.04 | 0.039381
| Mean |  | 0.040533 | 0.039836 | 0.037333 | 0.04135 | 0.041 | 0.039667 | 0.038833 | 0.041333 | 0.040167 | 0.041408 | 0.040183 | 0.0415 | 0.040167 | 0.038633 | 0.036833 | 0.02875 | 0.043333 | 0.04 | 0.049288
| Median |  | 0.04065 | 0.03997 | 0.0365 | 0.04145 | 0.041 | 0.04 | 0.039 | 0.0415 | 0.04 | 0.041637 | 0.0404 | 0.0415 | 0.04 | 0.0392 | 0.037 | 0.0279 | 0.04 | 0.04 | 0.04052
| Std Dev. |  | 0.00044121 | 0.00073327 | 0.0032042 | 0.00055767 | 0.00063246 | 0.0005164 | 0.00075277 | 0.0013663 | 0.00040825 | 0.0013517 | 0.00083526 | 0.00054772 | 0.001169 | 0.0013545 | 0.00075277 | 0.0029965 | 0.005164 | 0 | 0.022216
| Rel.Std.Dev. |  | 1.089% | 1.841% | 8.583% | 1.349% | 1.543% | 1.302% | 1.938% | 3.305% | 1.016% | 3.264% | 2.079% | 1.32% | 2.91% | 3.506% | 2.044% | 10.42% | 11.92% | 0% | 45.07%
| PDM3 |  | 1.475% | -0.2712% | -6.537% | 3.519% | 2.643% | -0.6951% | -2.781% | 3.477% | 0.5566% | 3.664% | 0.5983% | 3.895% | 0.5566% | -3.282% | -7.788% | -28.02% | 8.484% | 0.1394% | 23.39%
| Z-Score (Absolute) |  | 0.21 | 0.19 | 1.61 | 0.67 | 0.47 | 0.28 | 0.76 | 0.66 | 0 | 0.71 | 0.01 | 0.76 | 0 | 0.87 | 1.9 | 6.51 | 1.8 | 0.09 | 5.2
| Analytical results for Pb in OREAS 242b (Certified Value 19 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 19.5 | 19.0072 | 13 | 21 | 18.2 | 19.5 | 19.2 | 17.7 | 20.4 | 18.7254 | 20 | 19 | 17.48 | 21.1 | 18.01 | 18.4 | 17.68 | 18 | 19.737 |  | 1
| 1 | 2 | 19.8 | 18.9315 | 14 | 19 | 18.5 | 20 | 18.8 | 19.2 | 20.9 | 18.9351 | 16 | 19 | 16.99 | 19.2 | 18.42 | 18 | 19.19 | 18 | 20.849 |  | 16
| 1 | 3 | 19.5 | 19.0159 | 14 | 20 | 17.8 | 20.1 | 18.8 | 19.1 | 19.8 | 18.553 | 16 | 19 | 17.72 | 21.4 | 18.3 | 19.1 | 19.39 | 18 | 11.686 |  | 16
| 1 | 4 | 19.2 | 18.9008 | 16 | 19 | 17.3 | 19.6 | 18.8 | 19.3 | 19.9 | 19.037 | 18 | 19 | 17.07 | 19.2 | 17.97 | 19.1 | 18.91 | 18 | 20.513 |  | 19.0231
| 1 | 5 | 18.9 | 19.0914 | 13 | 20 | 17.4 | 19.8 | 18 | 18 | 19.2 | 19.0763 | 21 | 19 | 16.53 | 20.1 | 17.88 | 18.4 | 18.49 | 18 | 21.371 |  | 50
| 1 | 6 | 19.8 | 18.6427 | 17 | 19 | 18.1 | 19.9 | 19 | 17.8 | 19.4 | 18.4324 | 22 | 19 | 16.38 | 19.3 | 18.26 | 19.1 | 18.54 | 18 | 21.695
| Mean |  | 19.45 | 18.9316 | 14.5 | 19.6667 | 17.8833 | 19.8167 | 18.7667 | 18.5167 | 19.9333 | 18.7932 | 18.8333 | 19 | 17.0283 | 20.05 | 18.14 | 18.6833 | 18.7 | 18 | 19.3085
| Median |  | 19.5 | 18.9694 | 14 | 19.5 | 17.95 | 19.85 | 18.8 | 18.55 | 19.85 | 18.8303 | 19 | 19 | 17.03 | 19.7 | 18.135 | 18.75 | 18.725 | 18 | 20.681
| Std Dev. |  | 0.35071 | 0.15667 | 1.6432 | 0.8165 | 0.47081 | 0.23166 | 0.40825 | 0.75741 | 0.6314 | 0.2654 | 2.5626 | 0 | 0.52044 | 0.99348 | 0.21531 | 0.47924 | 0.61162 | 0 | 3.7963
| Rel.Std.Dev. |  | 1.803% | 0.8276% | 11.33% | 4.152% | 2.633% | 1.169% | 2.175% | 4.09% | 3.168% | 1.412% | 13.61% | 0% | 3.056% | 4.955% | 1.187% | 2.565% | 3.271% | 0% | 19.66%
| PDM3 |  | 2.244% | -0.4809% | -23.78% | 3.383% | -5.991% | 4.172% | -1.348% | -2.662% | 4.785% | -1.208% | -0.9974% | -0.1213% | -10.49% | 5.398% | -4.642% | -1.786% | -1.698% | -5.378% | 1.5%
| Z-Score (Absolute) |  | 0.68 | 0.14 | NA | 0.9 | 0.94 | 1.06 | 0.03 | 0.29 | 1.18 | 0 | NA | NA | 1.82 | 1.3 | 0.67 | 0.11 | 0.1 | NA | 0.53
Results from laboratories 1.07, 1.15, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 242b (Certified Value 2.21 ppm) | [21 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.26 | 2.3443 | 2.7715 | 2.4 | 2.3 | 2.18 | 2 | 2.16 | 2.15 | [11 empty] | 1
| 1 | 2 | 2.21 | 2.1416 | 2.6515 | 2.3 | 2.3 | 2.19 | 2.1 | 2.1 | 2.2 | [11 empty] | 33
| 1 | 3 | 2.25 | 2.3287 | 2.6636 | 2.3 | 2.3 | 2.23 | 2.1 | 2.13 | 2.2 | [11 empty] | 16
| 1 | 4 | 2.27 | 2.1922 | 2.7569 | 2.3 | 2.3 | 2.19 | 2 | 2.21 | 2.1 | [11 empty] | 2.2059
| 1 | 5 | 2.23 | 2.2349 | 2.8307 | 2.3 | 2.3 | 2.16 | 2 | 2.14 | 2.15 | [11 empty] | 51
| 1 | 6 | 2.27 | 2.2432 | 2.6467 | 2.3 | 2.3 | 2.14 | 2 | 2.23 | 2.15
| Mean |  | 2.2483 | 2.2475 | 2.7201 | 2.3167 | 2.3 | 2.1817 | 2.0333 | 2.1617 | 2.1583
| Median |  | 2.255 | 2.2391 | 2.7102 | 2.3 | 2.3 | 2.185 | 2 | 2.15 | 2.15
| Std Dev. |  | 0.024014 | 0.07798 | 0.076829 | 0.040825 | 4.8648e-16 | 0.030605 | 0.05164 | 0.049565 | 0.037639
| Rel.Std.Dev. |  | 1.068% | 3.47% | 2.824% | 1.762% | 2.115e-14% | 1.403% | 2.54% | 2.293% | 1.744%
| PDM3 |  | 1.922% | 1.884% | 23.31% | 5.02% | 4.264% | -1.1% | -7.825% | -2.007% | -2.158%
| Z-Score (Absolute) |  | 0.01 | 0 | 4.61 | 0.67 | 0.51 | 0.64 | 2.09 | 0.84 | 0.87
| Analytical results for Pt in OREAS 242b (Indicative Value 13.8 ppb) | [21 empty] | Indicative
Pt | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.21 | [19 empty] | ppb
|  |  | 4A*OES/MS | [19 empty] | 1
1
| 1 | 1 | 16 | [19 empty] | 1
| 1 | 2 | 14 | [19 empty] | 1
| 1 | 3 | 14 | [19 empty] | 16
| 1 | 4 | 12 | [19 empty] | 13.8333
| 1 | 5 | 14 | [19 empty] | 7
| 1 | 6 | 13
| Mean |  | 13.8333
| Median |  | 14
| Std Dev. |  | 1.3292
| Rel.Std.Dev. |  | 9.608%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 242b (Certified Value 14.7 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.86 | 14.1846 | 16.5642 | 14.1 | 14.8 | 14.8 | 15 | 14.6 | 15.2 | 14.0724 | 15.7 | 15.7 | 14.2 | 13.73 | 13.6 | 15 | 13.2 |  |  |  | 1
| 1 | 2 | 14.93 | 14.0215 | 16.6686 | 14 | 14.7 | 14.9 | 14.8 | 15.6 | 15.7 | 14.7145 | 16.5 | 15.9 | 13.4 | 13.71 | 14 | 14.2 | 13.2 |  |  |  | 34
| 1 | 3 | 15.37 | 14.7337 | 16.6087 | 13.8 | 14.2 | 15 | 15.2 | 15.6 | 15.4 | 14.3408 | 15.4 | 15.4 | 13.6 | 14.16 | 13.8 | 14.8 | 13.2 |  |  |  | 16
| 1 | 4 | 15.29 | 14.3472 | 16.6979 | 13.9 | 14 | 14.6 | 14.9 | 15.3 | 15.3 | 14.708 | 15.1 | 15.2 | 13.9 | 14.36 | 13.9 | 15.2 | 13 |  |  |  | 14.6888
| 1 | 5 | 15.3 | 14.079 | 16.6193 | 13.4 | 14.4 | 14.9 | 14.4 | 14.6 | 14.8 | 14.5702 | 15.7 | 15.8 | 13.5 | 13.85 | 13.7 | 14.6 | 13.8 |  |  |  | 52
| 1 | 6 | 15.06 | 14.6726 | 16.6449 | 13.9 | 14.5 | 15.2 | 14.8 | 14.6 | 15.6 | 14.0764 | 15.6 | 15.9 | 13.4 | 13.91 | 14 | 14.7 | 13.2
| Mean |  | 15.135 | 14.3397 | 16.634 | 13.85 | 14.4333 | 14.9 | 14.85 | 15.05 | 15.3333 | 14.4137 | 15.6667 | 15.65 | 13.6667 | 13.9533 | 13.8333 | 14.75 | 13.2667
| Median |  | 15.175 | 14.2659 | 16.6321 | 13.9 | 14.45 | 14.9 | 14.85 | 14.95 | 15.35 | 14.4555 | 15.65 | 15.75 | 13.55 | 13.88 | 13.85 | 14.75 | 13.2
| Std Dev. |  | 0.21436 | 0.30313 | 0.047168 | 0.2429 | 0.30111 | 0.2 | 0.26646 | 0.50498 | 0.32042 | 0.29568 | 0.46762 | 0.2881 | 0.32042 | 0.2568 | 0.1633 | 0.34496 | 0.27325
| Rel.Std.Dev. |  | 1.416% | 2.114% | 0.2836% | 1.754% | 2.086% | 1.342% | 1.794% | 3.355% | 2.09% | 2.051% | 2.985% | 1.841% | 2.345% | 1.84% | 1.18% | 2.339% | 2.06%
| PDM3 |  | 3.038% | -2.376% | 13.24% | -5.71% | -1.739% | 1.438% | 1.098% | 2.459% | 4.388% | -1.873% | 6.657% | 6.544% | -6.958% | -5.007% | -5.824% | 0.4169% | -9.682%
| Z-Score (Absolute) |  | 0.45 | 0.47 | 2.18 | 1.04 | 0.37 | 0.17 | 0.12 | 0.35 | 0.67 | 0.39 | 1.06 | 1.04 | 1.25 | 0.92 | 1.06 | 0 | 1.71
| Analytical results for Re in OREAS 242b (Certified Value 0.002 ppm) | [21 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.003 | 0.0094981 | 0.002 | < 0.002 | 0.002 | 0.002 | 0.003 | 0.003 | < 0.05 | 0.009 | < 0.05 | 0.003 | 0.002 | < 0.002 | 0.0012 | < 0.1 |  |  |  |  | 1
| 1 | 2 |  | 0.010694 | 0.002 | 0.002 | 0.003 | 0.002 | 0.003 | 0.003 | < 0.05 | 0.007 | < 0.05 | 0.002 |  | < 0.002 | 0.001 | < 0.1 |  |  |  |  | 35
| 1 | 3 |  | 0.010081 | 0.002 | < 0.002 | 0.002 | 0.003 | 0.002 | 0.002 | < 0.05 | 0.008 | < 0.05 | < 0.002 |  | < 0.002 | 0.001 | < 0.1 |  |  |  |  | 16
| 1 | 4 |  | 0.011319 | 0.002 | 0.002 | 0.002 | < 0.002 | 0.003 | 0.003 | < 0.05 | 0.008 | < 0.05 | 0.003 |  | < 0.002 | 0.0012 | < 0.1 |  |  |  |  | 0.0024556
| 1 | 5 |  | 0.010484 | 0.002 | 0.002 | < 0.002 | 0.003 | 0.003 | 0.003 | < 0.05 | 0.009 | < 0.05 | < 0.002 |  | < 0.002 | 0.0008 | < 0.1 |  |  |  |  | 53
| 1 | 6 | 0.002 | 0.011129 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.003 | < 0.05 | 0.009 | < 0.05 | 0.004 | 0.003 | < 0.002 | 0.0011 | < 0.1
| Mean |  | 0.0025 | 0.010534 | 0.002 | 0.002 | 0.0022 | 0.0024 | 0.0026667 | 0.0028333 |  | 0.0083333 |  | 0.003 | 0.0025 |  | 0.00105
| Median |  | 0.0025 | 0.010589 | 0.002 | 0.002 | 0.002 | 0.002 | 0.003 | 0.003 |  | 0.0085 |  | 0.003 | 0.0025 |  | 0.00105
| Std Dev. |  | 0.00070711 | 0.00067488 | 0 | 0 | 0.00044721 | 0.00054772 | 0.0005164 | 0.00040825 |  | 0.0008165 |  | 0.0008165 | 0.00070711 |  | 0.00015166
| Rel.Std.Dev. |  | 28.28% | 6.407% | 0% | 0% | 20.33% | 22.82% | 19.36% | 14.41% |  | 9.798% |  | 27.22% | 28.28% |  | 14.44%
| PDM3 |  | 1.81% | 329% | -18.55% | -18.55% | -10.41% | -2.262% | 8.597% | 15.38% |  | 239.4% |  | 22.17% | 1.81% |  | -57.24%
| Z-Score (Absolute) |  | 0 | 8.13 | 0.51 | 0.84 | 0.51 | 0.34 | 0.17 | 0.34 | 22.76 | 5.9 | 22.76 | 0.17 | 0 | 1.52 | 1.47 | 48.04
| Analytical results for S in OREAS 242b (Certified Value 0.265 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.2714 | 0.26944 | 0.28529 | 0.25 | 0.29 | 0.26 | 0.26 | 0.28 | 0.26 | 0.27 | 0.26679 | 0.2522 | 0.285 | 0.27 | 0.24 | 0.24 | 0.254 | 0.14 | 0.27 | 0.31784 | 1
| 1 | 2 | 0.2676 | 0.26659 | 0.28372 | 0.25 | 0.28 | 0.26 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27128 | 0.2616 | 0.292 | 0.27 | 0.23 | 0.24 | 0.251 | 0.14 | 0.265 | 0.31754 | 20
| 1 | 3 | 0.2711 | 0.28173 | 0.28298 | 0.25 | 0.29 | 0.26 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27584 | 0.2618 | 0.289 | 0.25 | 0.24 | 0.24 | 0.238 | 0.15 | 0.265 | 0.017667 | 16
| 1 | 4 | 0.2704 | 0.27081 | 0.2828 | 0.26 | 0.28 | 0.26 | 0.26 | 0.26 | 0.27 | 0.27 | 0.27925 | 0.2532 | 0.292 | 0.27 | 0.25 | 0.24 | 0.244 | 0.14 | 0.265 | 0.32037 | 0.26461
| 1 | 5 | 0.2762 | 0.26677 | 0.28283 | 0.26 | 0.28 | 0.26 | 0.28 | 0.26 | 0.27 | 0.27 | 0.2699 | 0.252 | 0.291 | 0.26 | 0.22 | 0.24 | 0.23 | 0.13 | 0.27 | 0.31085 | 54
| 1 | 6 | 0.2681 | 0.27506 | 0.28429 | 0.26 | 0.28 | 0.26 | 0.26 | 0.27 | 0.27 | 0.26 | 0.27759 | 0.2594 | 0.284 | 0.26 | 0.23 | 0.24 | 0.235 | 0.14 | 0.27 | 0.30961
| Mean |  | 0.2708 | 0.27173 | 0.28365 | 0.255 | 0.28333 | 0.26 | 0.26667 | 0.26833 | 0.26833 | 0.26833 | 0.27344 | 0.2567 | 0.28883 | 0.26333 | 0.235 | 0.24 | 0.242 | 0.14 | 0.2675 | 0.26564
| Median |  | 0.27075 | 0.27012 | 0.28335 | 0.255 | 0.28 | 0.26 | 0.265 | 0.27 | 0.27 | 0.27 | 0.27356 | 0.2563 | 0.29 | 0.265 | 0.235 | 0.24 | 0.241 | 0.14 | 0.2675 | 0.31419
| Std Dev. |  | 0.0030757 | 0.0058005 | 0.00099402 | 0.0054772 | 0.005164 | 0 | 0.008165 | 0.0075277 | 0.0040825 | 0.0040825 | 0.0048607 | 0.0047308 | 0.0035449 | 0.008165 | 0.010488 | 0 | 0.0093595 | 0.0063246 | 0.0027386 | 0.12156
| Rel.Std.Dev. |  | 1.136% | 2.135% | 0.3504% | 2.148% | 1.823% | 0% | 3.062% | 2.805% | 1.521% | 1.521% | 1.778% | 1.843% | 1.227% | 3.101% | 4.463% | 0% | 3.868% | 4.518% | 1.024% | 45.76%
| PDM3 |  | 2.339% | 2.692% | 7.196% | -3.632% | 7.076% | -1.742% | 0.777% | 1.407% | 1.407% | 1.407% | 3.337% | -2.99% | 9.154% | -0.4827% | -11.19% | -9.301% | -8.545% | -47.09% | 1.092% | 0.3906%
| Z-Score (Absolute) |  | 0.37 | 0.47 | 1.66 | 1.21 | 1.63 | 0.71 | 0.04 | 0.13 | 0.13 | 0.13 | 0.64 | 1.04 | 2.18 | 0.38 | 3.22 | 2.72 | 2.52 | 12.75 | 0.04 | 0.14
| Analytical results for Sb in OREAS 242b (Certified Value 3.46 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 3.47 | 3.1503 | 3.6 | < 5 | 3.58 | 3.56 | 3.62 | 3.43 | 3.76 | 3.596 | 3.24 | 3.3 | 3.1 | 3.54 | 2.98 | 3.4 | < 5 | 4 | 3.696 |  | 1
| 1 | 2 | 3.45 | 3.0708 | 3.6 | < 5 | 3.69 | 3.44 | 3.49 | 3.7 | 3.93 | 3.6249 | 3.21 | 3.3 | 3.04 | 3.61 | 2.98 | 3.3 | < 5 | 3.7 | 3.258 |  | 36
| 1 | 3 | 3.4 | 3.2104 | 3.3 | < 5 | 3.35 | 3.37 | 3.5 | 3.63 | 3.76 | 3.6025 | 3.16 | 3.3 | 3.34 | 3.57 | 2.96 | 3.6 | < 5 | 3.8 | 4.31 |  | 16
| 1 | 4 | 3.48 | 3.0973 | 3.6 | < 5 | 3.36 | 3.35 | 3.61 | 3.72 | 3.81 | 3.831 | 3.09 | 3.3 | 3.23 | 3.56 | 2.96 | 3.5 | < 5 | 3.7 | 1.505 |  | 3.4569
| 1 | 5 | 3.42 | 3.3449 | 3.6 | < 5 | 3.59 | 3.39 | 3.28 | 3.62 | 3.7 | 3.7986 | 3.21 | 3.3 | 2.94 | 3.61 | 3 | 3.5 | < 5 | 4 | 4.575 |  | 55
| 1 | 6 | 3.53 | 3.1539 | 3.2 | < 5 | 3.47 | 3.46 | 3.43 | 3.49 | 3.78 | 3.4678 | 3.18 | 3.3 | 2.93 | 3.57 | 2.95 | 3.6 | < 5 | 3.8 | 3.606
| Mean |  | 3.4583 | 3.1713 | 3.4833 |  | 3.5067 | 3.4283 | 3.4883 | 3.5983 | 3.79 | 3.6535 | 3.1817 | 3.3 | 3.0967 | 3.5767 | 2.9717 | 3.4833 |  | 3.8333 | 3.4917
| Median |  | 3.46 | 3.1521 | 3.6 |  | 3.525 | 3.415 | 3.495 | 3.625 | 3.77 | 3.6137 | 3.195 | 3.3 | 3.07 | 3.57 | 2.97 | 3.5 |  | 3.8 | 3.651
| Std Dev. |  | 0.046224 | 0.097957 | 0.18348 |  | 0.13663 | 0.07679 | 0.12576 | 0.11548 | 0.07746 | 0.13695 | 0.052694 | 4.8648e-16 | 0.16281 | 0.028048 | 0.018348 | 0.1169 |  | 0.13663 | 1.0862
| Rel.Std.Dev. |  | 1.337% | 3.089% | 5.268% |  | 3.896% | 2.24% | 3.605% | 3.209% | 2.044% | 3.749% | 1.656% | 1.474e-14% | 5.258% | 0.7842% | 0.6174% | 3.356% |  | 3.564% | 31.11%
| PDM3 |  | 0.04101% | -8.263% | 0.7642% |  | 1.439% | -0.8268% | 0.9088% | 4.091% | 9.635% | 5.686% | -7.962% | -4.539% | -10.42% | 3.464% | -14.04% | 0.7642% |  | 10.89% | 1.005%
| Z-Score (Absolute) |  | 0.1 | 1.24 | 0 | 3.9 | 0.09 | 0.22 | 0.02 | 0.46 | 1.22 | 0.67 | 1.2 | 0.73 | 1.53 | 0.37 | 2.03 | 0 | 3.9 | 1.39 | 0.03
| Analytical results for Sc in OREAS 242b (Certified Value 37.2 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 38.2 | 37.501 | 37.894 | 37 | 42.6 | 38.3 | 37.6 | 38.9 | 36.8 | 34.9 | 38.425 | 34.53 | 38 | 37.1 | 34.2 | 34.5 | 43 | 37 | 35.665 |  | 1
| 1 | 2 | 38.7 | 36.7616 | 37.925 | 37 | 41.3 | 37.9 | 37.5 | 38.7 | 40.7 | 37.1 | 40.2694 | 35.1 | 38 | 37.9 | 34.3 | 35.5 | 42 | 37 | 35.587 |  | 37
| 1 | 3 | 39.4 | 38.5638 | 37.936 | 37 | 40.2 | 36.4 | 37.4 | 39.4 | 38.5 | 34.7 | 37.9386 | 34.56 | 37 | 35.6 | 34.4 | 34.8 | 41 | 36 | 12.549 |  | 16
| 1 | 4 | 38.4 | 37.6394 | 37.976 | 37 | 41.6 | 36.3 | 36.9 | 38.5 | 39.1 | 34.3 | 40.6198 | 34.47 | 39 | 38.5 | 34.1 | 36.2 | 44 | 37 | 36.212 |  | 37.1853
| 1 | 5 | 38.6 | 36.8992 | 37.932 | 37 | 40.7 | 37.6 | 37.2 | 37.4 | 37 | 33.4 | 39.5204 | 34.22 | 38 | 37.3 | 34.8 | 35.5 | 39 | 38 | 35.103 |  | 56
| 1 | 6 | 38.9 | 37.8368 | 37.887 | 37 | 38.4 | 37.6 | 38.4 | 39.1 | 37.1 | 36 | 37.6552 | 34.81 | 37 | 35.7 | 33.7 | 35.2 | 42 | 37 | 35.054
| Mean |  | 38.7 | 37.5336 | 37.925 | 37 | 40.8 | 37.35 | 37.5 | 38.6667 | 38.2 | 35.0667 | 39.0714 | 34.615 | 37.8333 | 37.0167 | 34.25 | 35.2833 | 41.8333 | 37 | 31.695
| Median |  | 38.65 | 37.5702 | 37.9285 | 37 | 41 | 37.6 | 37.45 | 38.8 | 37.8 | 34.8 | 38.9727 | 34.545 | 38 | 37.2 | 34.25 | 35.35 | 42 | 37 | 35.345
| Std Dev. |  | 0.41952 | 0.65787 | 0.032175 | 0 | 1.4325 | 0.8167 | 0.50596 | 0.69474 | 1.5336 | 1.3064 | 1.2443 | 0.3035 | 0.75277 | 1.1669 | 0.36194 | 0.59805 | 1.7224 | 0.63246 | 9.3891
| Rel.Std.Dev. |  | 1.084% | 1.753% | 0.08484% | 0% | 3.511% | 2.187% | 1.349% | 1.797% | 4.015% | 3.725% | 3.185% | 0.8768% | 1.99% | 3.152% | 1.057% | 1.695% | 4.117% | 1.709% | 29.62%
| PDM3 |  | 4.073% | 0.9367% | 1.989% | -0.4984% | 9.721% | 0.4428% | 0.8462% | 3.984% | 2.729% | -5.698% | 5.072% | -6.912% | 1.743% | -0.4536% | -7.894% | -5.115% | 12.5% | -0.4984% | -14.76%
| Z-Score (Absolute) |  | 0.69 | 0.02 | 0.25 | 0.29 | 1.91 | 0.09 | 0 | 0.67 | 0.4 | 1.41 | 0.91 | 1.67 | 0.19 | 0.28 | 1.88 | 1.28 | 2.5 | 0.29 | 3.36
| Analytical results for Se in OREAS 242b (Indicative Value 1.12 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  | 2
2
| 1 | 1 | 1.1 | < 1.5 | 0.5 | < 5 | 2 | 1 | 1 | 2 | 1 | < 2 | 0.9 | < 1 | 1.7 | 0.8 | 1.1 | 1.76 | < 5 | 3.292 |  |  | 1
| 1 | 2 | 1.1 | < 1.5 | 0.4 | < 5 | 1 | 1 | 1 | < 1 | 1 | < 2 | 0.91 | < 1 | 1.3 | 0.7 | 1 | 1.7 | < 5 | 2.644 |  |  | 2
| 1 | 3 | 0.9 | < 1.5 | 0.3 | < 5 | 2 | 1 | 1 | 2 | 1 | < 2 | 0.84 | < 1 | 0.9 | 0.9 | 1.2 | 1.51 | < 5 | 4.462 |  |  | 16
| 1 | 4 | 1 | < 1.5 | 0.6 | < 5 | 2 | 2 | 1 | 1 | 1 | < 2 | 0.87 | < 1 | 1.2 | 0.9 | 1.2 | 1.51 | < 5 | 4.254 |  |  | 1.1232
| 1 | 5 | 0.8 | < 1.5 | 0.9 | < 5 | 1 | 1 | 1 | 1 | 1 | < 2 | 0.87 | < 1 | 0.8 | 0.9 | 1.3 | 1.34 | < 5 | 3.541 |  |  | 8
| 1 | 6 | 0.8 | < 1.5 | 0.8 | < 5 | 1 | 2 | 1 | 1 | 1 | < 2 | 0.9 | < 1 | 1.7 | 0.7 | 1.1 | 1.76 | < 5 | 7.132
| Mean |  | 0.95 |  | 0.58333 |  | 1.5 | 1.3333 | 1 | 1.4 | 1 |  | 0.88167 |  | 1.2667 | 0.81667 | 1.15 | 1.5967 |  | 4.2208
| Median |  | 0.95 |  | 0.55 |  | 1.5 | 1 | 1 | 1 | 1 |  | 0.885 |  | 1.25 | 0.85 | 1.15 | 1.605 |  | 3.8975
| Std Dev. |  | 0.13784 |  | 0.23166 |  | 0.54772 | 0.5164 | 0 | 0.54772 | 0 |  | 0.026394 |  | 0.38297 | 0.098319 | 0.10488 | 0.17025 |  | 1.5713
| Rel.Std.Dev. |  | 14.51% |  | 39.71% |  | 36.51% | 38.73% | 0% | 39.12% | 0% |  | 2.994% |  | 30.23% | 12.04% | 9.12% | 10.66% |  | 37.23%
| PDM3 |  | -15.42% |  | -48.06% |  | 33.55% | 18.71% | -10.97% | 24.64% | -10.97% |  | -21.5% |  | 12.77% | -27.29% | 2.387% | 42.15% |  | 275.8%
| Z-Score (Absolute) |  | 0.33 | 0.85 | 1.28 | 3.72 | 1.11 | 0.67 | 0.2 | 0.46 | 0.2 | 0.2 | 0.5 | 1.5 | 0.5 | 0.67 | 0.2 | 1.36 | 3.72 | 8.21
| Analytical results for Sm in OREAS 242b (Certified Value 2.78 ppm) | [21 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.8 | 2.9452 | 3.0356 | 3.1 | 3.1 | 2.74 | 2.5 | 2.68 | 2.7 | [11 empty] | 1
| 1 | 2 | 2.77 | 2.7094 | 3.0825 | 2.8 | 2.9 | 2.69 | 2.6 | 2.63 | 2.7 | [11 empty] | 7
| 1 | 3 | 2.89 | 2.8571 | 3.016 | 2.9 | 2.7 | 2.69 | 2.6 | 2.74 | 2.7 | [11 empty] | 16
| 1 | 4 | 2.9 | 2.7969 | 3.0236 | 3.1 | 2.9 | 2.56 | 2.6 | 2.55 | 2.7 | [11 empty] | 2.7846
| 1 | 5 | 2.75 | 2.7681 | 3.0728 | 2.7 | 2.7 | 2.7 | 2.5 | 2.79 | 2.8 | [11 empty] | 57
| 1 | 6 | 2.87 | 2.6501 | 2.9942 | 2.8 | 2.7 | 2.71 | 2.6 | 2.81 | 2.7
| Mean |  | 2.83 | 2.7878 | 3.0375 | 2.9 | 2.8333 | 2.6817 | 2.5667 | 2.7 | 2.7167
| Median |  | 2.835 | 2.7825 | 3.0296 | 2.85 | 2.8 | 2.695 | 2.6 | 2.71 | 2.7
| Std Dev. |  | 0.064807 | 0.10501 | 0.034062 | 0.16733 | 0.1633 | 0.062423 | 0.05164 | 0.099599 | 0.040825
| Rel.Std.Dev. |  | 2.29% | 3.767% | 1.121% | 5.77% | 5.764% | 2.328% | 2.012% | 3.689% | 1.503%
| PDM3 |  | 1.631% | 0.1158% | 9.081% | 4.145% | 1.751% | -3.696% | -7.826% | -3.038% | -2.439%
| Z-Score (Absolute) |  | 0.32 | 0 | 1.92 | 0.86 | 0.35 | 0.82 | 1.7 | 0.67 | 0.55
| Analytical results for Sn in OREAS 242b (Certified Value 1 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
2
| 1 | 1 | 1.1 | 1.0242 | 1 | 0.9 | 1 | 0.9 | 0.9 | 1 | 1.1 | 0.85358 | 0.8 | 1 | 1 | 1 | 0.8 | 0.2 | < 5 | 1.2 | 1.872 |  | 1
| 1 | 2 | 0.9 | 1.0332 | < 1 | 0.8 | 1 | 1 | 1 | 1 | 1.1 | 0.91921 | 1.2 | 1 | 1 | 1.1 | 0.8 | 0.3 | < 5 | 1.1 | 0.864 |  | 11
| 1 | 3 | 1 | 1.0027 | < 1 | 0.9 | 1 | 1.1 | 0.9 | 1 | 1.1 | 0.9282 | 1.4 | 1 | 1.9 | 1.1 | 0.8 | 0.4 | < 5 | 1.1 | 9.286 |  | 16
| 1 | 4 | 1 | 1.0178 | < 1 | 0.8 | 1 | 0.9 | 0.9 | 1 | 1.1 | 0.91375 | 0.8 | 1 | 1.2 | 1 | 0.9 | 0.3 | < 5 | 1.1 | 1.712 |  | 0.99669
| 1 | 5 | 1 | 0.98235 | 1 | 0.9 | 1 | 1 | 0.9 | 1 | 1.1 | 0.93563 | 1.5 | 1 | 0.9 | 1.1 | 0.8 | 0.3 | < 5 | 1.1 | 1.804 |  | 58
| 1 | 6 | 1 | 1.01 | < 1 | 0.9 | 1 | 1 | 0.9 | 1 | 1 | 0.84116 | 1.8 | 1 | 1.2 | 1 | 0.9 | 0.2 | < 5 | 1.1 | 1.077
| Mean |  | 1 | 1.0117 | 1 | 0.86667 | 1 | 0.98333 | 0.91667 | 1 | 1.0833 | 0.89859 | 1.25 | 1 | 1.2 | 1.05 | 0.83333 | 0.28333 |  | 1.1167 | 2.7692
| Median |  | 1 | 1.0139 | 1 | 0.9 | 1 | 1 | 0.9 | 1 | 1.1 | 0.91648 | 1.3 | 1 | 1.1 | 1.05 | 0.8 | 0.3 |  | 1.1 | 1.758
| Std Dev. |  | 0.063246 | 0.017896 | 0 | 0.05164 | 0 | 0.075277 | 0.040825 | 0 | 0.040825 | 0.040564 | 0.39875 | 0 | 0.36332 | 0.054772 | 0.05164 | 0.075277 |  | 0.040825 | 3.2192
| Rel.Std.Dev. |  | 6.325% | 1.769% | 0% | 5.958% | 0% | 7.655% | 4.454% | 0% | 3.768% | 4.514% | 31.9% | 0% | 30.28% | 5.216% | 6.197% | 26.57% |  | 3.656% | 116.3%
| PDM3 |  | 0.3319% | 1.508% | 0.3319% | -13.05% | 0.3319% | -1.34% | -8.029% | 0.3319% | 8.693% | -9.843% | 25.41% | 0.3319% | 20.4% | 5.348% | -16.39% | -71.57% |  | 12.04% | 177.8%
| Z-Score (Absolute) |  | 0.07 | 0 | NA | 0.86 | NA | 0.17 | 0.57 | NA | 0.43 | 0.67 | 1.42 | NA | 1.12 | 0.23 | 1.06 | 4.34 | 8.87 | 0.63 | 10.48
Results from laboratories 1.07, 1.09, 1.12 and 1.16 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 242b (Certified Value 185 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 188.61 | 182.6427 | 188.2736 | 199 | 195 | 197.5 | 188.5 | 198 | 187 | 189.5 | 176.3985 | 180 | 184 | 190 | 176.25 | 172 | 184 | 184.5 | 185 | 174.376 | 1
| 1 | 2 | 187.08 | 182.288 | 187.6376 | 192 | 180 | 195 | 193 | 197.5 | 201 | 190.5 | 183.3395 | 180 | 188 | 191 | 177.69 | 178 | 172 | 181.7 | 184 | 176.475 | 12
| 1 | 3 | 194.24 | 187.8084 | 185.5706 | 192 | 185 | 193.5 | 192.5 | 194.5 | 197 | 190.5 | 178.9133 | 175 | 188 | 183 | 180.78 | 176 | 161 | 187 | 184 | 265.897 | 16
| 1 | 4 | 192.35 | 182.9239 | 185.7932 | 189 | 186 | 194.5 | 188 | 194.5 | 199 | 189 | 183.8924 | 172 | 186 | 194 | 177.48 | 178 | 174 | 183.2 | 181 | 175.056 | 185.2871
| 1 | 5 | 192.03 | 178.5099 | 186.3126 | 195 | 187 | 196 | 196 | 190 | 197.5 | 184.5 | 176.6018 | 175 | 185 | 186 | 180.37 | 172 | 156 | 182.3 | 183 | 173.268 | 59
| 1 | 6 | 193.31 | 184.646 | 186.037 | 190 | 187 | 196 | 193 | 195.5 | 198 | 187.5 | 181.1808 | 176 | 187 | 182 | 176.58 | 177 | 166 | 180.9 | 185 | 168.254
| Mean |  | 191.27 | 183.1365 | 186.6041 | 192.8333 | 186.6667 | 195.4167 | 191.8333 | 195 | 196.5833 | 188.5833 | 180.0544 | 176.3333 | 186.3333 | 187.6667 | 178.1917 | 175.5 | 168.8333 | 183.2667 | 183.6667 | 188.8877
| Median |  | 192.19 | 182.7833 | 186.1748 | 192 | 186.5 | 195.5 | 192.75 | 195 | 197.75 | 189.25 | 180.047 | 175.5 | 186.5 | 188 | 177.585 | 176.5 | 169 | 182.75 | 184 | 174.716
| Std Dev. |  | 2.8056 | 3.0502 | 1.0943 | 3.656 | 4.8442 | 1.3934 | 3.0441 | 2.8636 | 4.9032 | 2.2895 | 3.2674 | 3.1411 | 1.633 | 4.761 | 1.9272 | 2.8107 | 10.0083 | 2.2115 | 1.5055 | 37.8316
| Rel.Std.Dev. |  | 1.467% | 1.666% | 0.5864% | 1.896% | 2.595% | 0.7131% | 1.587% | 1.468% | 2.494% | 1.214% | 1.815% | 1.781% | 0.8764% | 2.537% | 1.082% | 1.602% | 5.928% | 1.207% | 0.8197% | 20.03%
| PDM3 |  | 3.229% | -1.161% | 0.7108% | 4.073% | 0.7445% | 5.467% | 3.533% | 5.242% | 6.097% | 1.779% | -2.824% | -4.832% | 0.5646% | 1.284% | -3.829% | -5.282% | -8.88% | -1.09% | -0.8746% | 1.943%
| Z-Score (Absolute) |  | 0.64 | 0.48 | 0 | 0.85 | 0 | 1.2 | 0.71 | 1.15 | 1.36 | 0.27 | 0.9 | 1.41 | 0.04 | 0.14 | 1.16 | 1.53 | 2.44 | 0.46 | 0.41 | 0.31
| Analytical results for Ta in OREAS 242b (Certified Value 0.26 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.28 | 0.26836 | < 0.1 | 10 | 0.26 | 0.25 | 0.26 | 0.25 | 0.26 | 0.25834 | 0.3 | 0.26 | 0.4 | 0.27 | 0.2 | 0.25 | 0.3 |  |  |  | 1
| 1 | 2 | 0.27 | 0.25633 | < 0.1 | 9 | 0.25 | 0.24 | 0.26 | 0.27 | 0.26 | 0.26499 | 0.31 | 0.28 | 0.4 | 0.26 | 0.2 | 0.28 | 0.3 |  |  |  | 25
| 1 | 3 | 0.28 | 0.26633 | < 0.1 | 9 | 0.25 | 0.24 | 0.26 | 0.27 | 0.26 | 0.25521 | 0.3 | 0.26 | 0.3 | 0.26 | 0.2 | 0.26 | 0.3 |  |  |  | 16
| 1 | 4 | 0.25 | 0.25036 | < 0.1 | 10 | 0.24 | 0.24 | 0.26 | 0.26 | 0.26 | 0.27743 | 0.3 | 0.25 | 0.3 | 0.27 | 0.2 | 0.27 | 0.3 |  |  |  | 0.26345
| 1 | 5 | 0.26 | 0.25045 | 0.3 | 9 | 0.25 | 0.24 | 0.26 | 0.28 | 0.27 | 0.2728 | 0.3 | 0.25 | 0.3 | 0.27 | 0.2 | 0.26 | 0.3 |  |  |  | 60
| 1 | 6 | 0.28 | 0.26701 | 0.3 | 9 | 0.24 | 0.25 | 0.26 | 0.25 | 0.27 | 0.26896 | 0.29 | 0.27 | 0.3 | 0.26 | 0.2 | 0.25 | 0.2
| Mean |  | 0.27 | 0.25981 | 0.3 | 9.3333 | 0.24833 | 0.24333 | 0.26 | 0.26333 | 0.26333 | 0.26629 | 0.3 | 0.26167 | 0.33333 | 0.265 | 0.2 | 0.26167 | 0.28333
| Median |  | 0.275 | 0.26133 | 0.3 | 9 | 0.25 | 0.24 | 0.26 | 0.265 | 0.26 | 0.26697 | 0.3 | 0.26 | 0.3 | 0.265 | 0.2 | 0.26 | 0.3
| Std Dev. |  | 0.012649 | 0.0084421 | 0 | 0.5164 | 0.0075277 | 0.005164 | 0 | 0.012111 | 0.005164 | 0.0084998 | 0.0063246 | 0.01169 | 0.05164 | 0.0054772 | 3.0405e-17 | 0.01169 | 0.040825
| Rel.Std.Dev. |  | 4.685% | 3.249% | 0% | 5.533% | 3.031% | 2.122% | 0% | 4.599% | 1.961% | 3.192% | 2.108% | 4.468% | 15.49% | 2.067% | 1.52e-14% | 4.468% | 14.41%
| PDM3 |  | 2.484% | -1.385% | 13.87% | 3443% | -5.74% | -7.638% | -1.311% | -0.04612% | -0.04612% | 1.075% | 13.87% | -0.6787% | 26.52% | 0.5865% | -24.09% | -0.6787% | 7.545%
| Z-Score (Absolute) |  | 1.35 | 0.71 | NA | 1834.79 | 3.03 | 4.05 | 0.67 | 0 | 0 | 0.6 | 7.42 | 0.34 | NA | 0.34 | NA | 0.34 | NA
Results from laboratories 1.07, 1.17, 1.19 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 242b (Certified Value 0.58 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
2
| 1 | 1 | 0.55 | 0.61051 | 0.71135 | 0.6 | 0.6 | 0.53 | 0.5 | 0.6 | 0.52 | 1.649 | [10 empty] | 1
| 1 | 2 | 0.52 | 0.57812 | 0.67789 | 0.6 | 0.6 | 0.53 | 0.5 | 0.55 | 0.52 | 1.513 | [10 empty] | 26
| 1 | 3 | 0.5 | 0.60991 | 0.69443 | 0.6 | 0.5 | 0.52 | 0.5 | 0.6 | 0.54 | 0.437 | [10 empty] | 16
| 1 | 4 | 0.56 | 0.57626 | 0.71603 | 0.6 | 0.6 | 0.51 | 0.5 | 0.61 | 0.54 | 2.256 | [10 empty] | 0.58011
| 1 | 5 | 0.54 | 0.56675 | 0.69695 | 0.6 | 0.5 | 0.53 | 0.5 | 0.59 | 0.54 | 1.814 | [10 empty] | 61
| 1 | 6 | 0.58 | 0.56034 | 0.69152 | 0.6 | 0.6 | 0.52 | 0.5 | 0.62 | 0.52 | 2.108
| Mean |  | 0.54167 | 0.58365 | 0.69803 | 0.6 | 0.56667 | 0.52333 | 0.5 | 0.595 | 0.53 | 1.6295
| Median |  | 0.545 | 0.57719 | 0.69569 | 0.6 | 0.6 | 0.525 | 0.5 | 0.6 | 0.53 | 1.7315
| Std Dev. |  | 0.028577 | 0.021566 | 0.013882 | 0 | 0.05164 | 0.008165 | 0 | 0.02429 | 0.010954 | 0.64676
| Rel.Std.Dev. |  | 5.276% | 3.695% | 1.989% | 0% | 9.113% | 1.56% | 0% | 4.082% | 2.067% | 39.69%
| PDM3 |  | -6.627% | 0.6093% | 20.33% | 3.428% | -2.318% | -9.788% | -13.81% | 2.566% | -8.638% | 180.9%
| Z-Score (Absolute) |  | 0.53 | 0 | 1.44 | NA | NA | 0.76 | NA | 0.14 | 0.67 | 13.15
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 242b (Certified Value 0.16 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | < 0.2 | 0.1 | 0.44 | 0.2 | 0.16 | 0.15 | 0.23 | 0.16 | 0.17965 | 0.15 | 0.1 | 0.15 | 0.16 | 0.14 | < 0.2 | [5 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.45 | 0.19 | 0.15 | 0.14 | 0.16 | 0.17 | 0.20755 | 0.12 | 0.1 | 0.15 | 0.19 | 0.14 | < 0.2 | [5 empty] | 27
| 1 | 3 | < 0.2 | < 0.1 | 0.44 | 0.17 | 0.16 | 0.15 | 0.22 | 0.16 | 0.17937 | 0.14 | 0.1 | 0.11 | 0.18 | 0.14 | < 0.2 | [5 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.46 | 0.17 | 0.16 | 0.13 | 0.22 | 0.16 | 0.21431 | 0.14 | 0.1 | 0.2 | 0.18 | 0.15 | < 0.2 | [5 empty] | 0.16479
| 1 | 5 | < 0.2 | 0.2 | 0.47 | 0.16 | 0.14 | 0.14 | 0.21 | 0.15 | 0.17962 | 0.12 | 0.1 | 0.11 | 0.17 | 0.15 | < 0.2 | [5 empty] | 62
| 1 | 6 | < 0.2 | 0.2 | 0.43 | 0.19 | 0.17 | 0.15 | 0.22 | 0.11 | 0.16706 | 0.17 | 0.1 | 0.13 | 0.17 | 0.14 | < 0.2
| Mean |  |  | 0.16667 | 0.44833 | 0.18 | 0.15667 | 0.14333 | 0.21 | 0.15167 | 0.18793 | 0.14 | 0.1 | 0.14167 | 0.175 | 0.14333
| Median |  |  | 0.2 | 0.445 | 0.18 | 0.16 | 0.145 | 0.22 | 0.16 | 0.17964 | 0.14 | 0.1 | 0.14 | 0.175 | 0.14
| Std Dev. |  |  | 0.057735 | 0.01472 | 0.015492 | 0.010328 | 0.008165 | 0.025298 | 0.02137 | 0.018588 | 0.018974 | 1.5202e-17 | 0.033714 | 0.010488 | 0.005164
| Rel.Std.Dev. |  |  | 34.64% | 3.283% | 8.607% | 6.592% | 5.696% | 12.05% | 14.09% | 9.891% | 13.55% | 1.52e-14% | 23.8% | 5.993% | 3.603%
| PDM3 |  |  | 1.137% | 172.1% | 9.228% | -4.931% | -13.02% | 27.43% | -7.965% | 14.04% | -15.04% | -39.32% | -14.03% | 6.194% | -13.02%
| Z-Score (Absolute) |  | 1.49 | NA | 8.57 | 0.82 | 0.14 | 0.24 | 1.69 | 0 | 1.05 | 0.34 | NA | 0.29 | 0.67 | 0.24 | 1.49
Results from laboratories 1.07 and 1.16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 242b (Certified Value 1.63 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.62 | 1.5199 | 1.8598 | 1.9 | 1.7 | 1.56 | 1.71 | 1.6 | 1.54 | 1.6198 | 1.64 | 1.69 | 1.5 | 1.62 | 1.5 | 1.42 | 1.6 |  |  |  | 1
| 1 | 2 | 1.62 | 1.4907 | 1.8556 | 1.8 | 1.76 | 1.58 | 1.68 | 1.72 | 1.63 | 1.6552 | 1.64 | 1.78 | 1.5 | 1.66 | 1.5 | 1.42 | 1.6 |  |  |  | 28
| 1 | 3 | 1.6 | 1.5597 | 1.7673 | 1.8 | 1.7 | 1.62 | 1.69 | 1.8 | 1.46 | 1.6284 | 1.6 | 1.74 | 1.5 | 1.7 | 1.5 | 1.52 | 1.7 |  |  |  | 16
| 1 | 4 | 1.59 | 1.4923 | 1.7677 | 1.9 | 1.64 | 1.57 | 1.68 | 1.72 | 1.49 | 1.7333 | 1.62 | 1.69 | 1.5 | 1.62 | 1.5 | 1.5 | 1.6 |  |  |  | 1.631
| 1 | 5 | 1.61 | 1.4876 | 1.8322 | 1.9 | 1.67 | 1.68 | 1.6 | 1.76 | 1.46 | 1.715 | 1.61 | 1.65 | 1.5 | 1.61 | 1.5 | 1.49 | 1.7 |  |  |  | 63
| 1 | 6 | 1.61 | 1.516 | 1.8368 | 1.9 | 1.7 | 1.6 | 1.68 | 1.58 | 1.5 | 1.645 | 1.61 | 1.7 | 1.5 | 1.68 | 1.5 | 1.49 | 1.6
| Mean |  | 1.6083 | 1.511 | 1.8199 | 1.8667 | 1.695 | 1.6017 | 1.6733 | 1.6967 | 1.5133 | 1.6661 | 1.62 | 1.7083 | 1.5 | 1.6483 | 1.5 | 1.4733 | 1.6333
| Median |  | 1.61 | 1.5041 | 1.8345 | 1.9 | 1.7 | 1.59 | 1.68 | 1.72 | 1.495 | 1.6501 | 1.615 | 1.695 | 1.5 | 1.64 | 1.5 | 1.49 | 1.6
| Std Dev. |  | 0.01169 | 0.02753 | 0.041931 | 0.05164 | 0.039875 | 0.044008 | 0.037771 | 0.088015 | 0.064395 | 0.046978 | 0.016733 | 0.04535 | 0 | 0.037103 | 0 | 0.04274 | 0.05164
| Rel.Std.Dev. |  | 0.7269% | 1.822% | 2.304% | 2.766% | 2.352% | 2.748% | 2.257% | 5.188% | 4.255% | 2.82% | 1.033% | 2.655% | 0% | 2.251% | 0% | 2.901% | 3.162%
| PDM3 |  | -1.389% | -7.354% | 11.58% | 14.45% | 3.925% | -1.798% | 2.597% | 4.027% | -7.213% | 2.154% | -0.6734% | 4.743% | -8.031% | 1.064% | -8.031% | -9.666% | 0.1441%
| Z-Score (Absolute) |  | 0.27 | 1.3 | 1.99 | 2.48 | 0.66 | 0.34 | 0.43 | 0.67 | 1.28 | 0.35 | 0.14 | 0.8 | 1.42 | 0.16 | 1.42 | 1.7 | 0
| Analytical results for Ti in OREAS 242b (Certified Value 0.568 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.5924 | 0.58188 | 0.57286 | 0.61 | 0.63 | 0.562 | 0.56 | 0.607 | 0.54 | 0.568 | 0.61672 | 0.6303 | 0.628 | 0.552 | 0.5605 | 0.559 | 0.581 | 0.53 | 0.56 | 0.56259 | 1
| 1 | 2 | 0.5811 | 0.57699 | 0.57333 | 0.597 | 0.62 | 0.554 | 0.578 | 0.584 | 0.584 | 0.574 | 0.62882 | 0.6468 | 0.645 | 0.566 | 0.56 | 0.583 | 0.545 | 0.52 | 0.565 | 0.56674 | 29
| 1 | 3 | 0.583 | 0.58949 | 0.57275 | 0.616 | 0.62 | 0.555 | 0.565 | 0.586 | 0.569 | 0.574 | 0.62241 | 0.6381 | 0.63 | 0.534 | 0.5672 | 0.575 | 0.513 | 0.53 | 0.57 | 0.60462 | 16
| 1 | 4 | 0.5844 | 0.58615 | 0.57244 | 0.598 | 0.63 | 0.558 | 0.561 | 0.579 | 0.568 | 0.563 | 0.63849 | 0.6298 | 0.635 | 0.569 | 0.5574 | 0.575 | 0.547 | 0.53 | 0.57 | 0.57443 | 0.56771
| 1 | 5 | 0.598 | 0.57669 | 0.57296 | 0.612 | 0.63 | 0.555 | 0.571 | 0.563 | 0.568 | 0.558 | 0.61901 | 0.6358 | 0.636 | 0.555 | 0.561 | 0.571 | 0.494 | 0.51 | 0.575 | 0.55861 | 64
| 1 | 6 | 0.5942 | 0.58328 | 0.57203 | 0.595 | 0.62 | 0.566 | 0.571 | 0.585 | 0.576 | 0.555 | 0.63463 | 0.6397 | 0.629 | 0.537 | 0.5613 | 0.575 | 0.536 | 0.53 | 0.56 | 0.55752
| Mean |  | 0.58885 | 0.58241 | 0.57273 | 0.60467 | 0.625 | 0.55833 | 0.56767 | 0.584 | 0.5675 | 0.56533 | 0.62668 | 0.63675 | 0.63383 | 0.55217 | 0.56123 | 0.573 | 0.536 | 0.525 | 0.56667 | 0.57075
| Median |  | 0.5884 | 0.58258 | 0.57281 | 0.604 | 0.625 | 0.5565 | 0.568 | 0.5845 | 0.5685 | 0.5655 | 0.62562 | 0.63695 | 0.6325 | 0.5535 | 0.56075 | 0.575 | 0.5405 | 0.53 | 0.5675 | 0.56467
| Std Dev. |  | 0.0069143 | 0.0050448 | 0.0004493 | 0.0090259 | 0.0054772 | 0.004761 | 0.0069186 | 0.014142 | 0.014856 | 0.0080416 | 0.008757 | 0.0063614 | 0.0063692 | 0.014442 | 0.0032377 | 0.0078994 | 0.030067 | 0.0083666 | 0.0060553 | 0.017696
| Rel.Std.Dev. |  | 1.174% | 0.8662% | 0.07845% | 1.493% | 0.8764% | 0.8527% | 1.219% | 2.422% | 2.618% | 1.422% | 1.397% | 0.999% | 1.005% | 2.615% | 0.5769% | 1.379% | 5.609% | 1.594% | 1.069% | 3.101%
| PDM3 |  | 3.723% | 2.589% | 0.8832% | 6.509% | 10.09% | -1.652% | -0.008458% | 2.869% | -0.03782% | -0.4195% | 10.39% | 12.16% | 11.65% | -2.739% | -1.142% | 0.931% | -5.586% | -7.524% | -0.1846% | 0.535%
| Z-Score (Absolute) |  | 0.9 | 0.56 | 0.05 | 1.73 | 2.8 | 0.7 | 0.21 | 0.64 | 0.22 | 0.34 | 2.89 | 3.42 | 3.26 | 1.03 | 0.55 | 0.07 | 1.88 | 2.46 | 0.27 | 0.05
| Analytical results for Tl in OREAS 242b (Certified Value 0.17 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 0.17 | 0.16997 | 0.17 | 0.16 | 0.18 | 0.2 | 0.17 | 0.18 | 0.16934 | 0.159 | 0.17 | 0.18 | 0.18 | 0.18 | 0.16 | 0.2 |  |  |  |  | 1
| 1 | 2 | 0.18 | 0.17286 | 0.17 | 0.17 | 0.19 | 0.2 | 0.18 | 0.2 | 0.18048 | 0.161 | 0.18 | 0.16 | 0.19 | 0.17 | 0.17 | 0.1 |  |  |  |  | 30
| 1 | 3 | 0.17 | 0.1708 | 0.16 | 0.16 | 0.19 | 0.2 | 0.17 | 0.18 | 0.16946 | 0.153 | 0.17 | 0.16 | 0.18 | 0.18 | 0.17 | 0.2 |  |  |  |  | 16
| 1 | 4 | 0.18 | 0.16115 | 0.17 | 0.16 | 0.18 | 0.18 | 0.17 | 0.18 | 0.17754 | 0.158 | 0.17 | 0.17 | 0.17 | 0.17 | 0.18 | 0.2 | 0.932 |  |  |  | 0.17172
| 1 | 5 | 0.17 | 0.16764 | 0.17 | 0.16 | 0.17 | 0.18 | 0.17 | 0.18 | 0.17362 | 0.155 | 0.17 | 0.18 | 0.18 | 0.17 | 0.17 | 0.2 |  |  |  |  | 65
| 1 | 6 | 0.17 | 0.16417 | 0.18 | 0.16 | 0.19 | 0.19 | 0.16 | 0.18 | 0.18048 | 0.161 | 0.17 | 0.16 | 0.18 | 0.18 | 0.17 | 0.2
| Mean |  | 0.17333 | 0.16777 | 0.17 | 0.16167 | 0.18333 | 0.19167 | 0.17 | 0.18333 | 0.17515 | 0.15783 | 0.17167 | 0.16833 | 0.18 | 0.175 | 0.17 | 0.18333 | 0.932
| Median |  | 0.17 | 0.16881 | 0.17 | 0.16 | 0.185 | 0.195 | 0.17 | 0.18 | 0.17558 | 0.1585 | 0.17 | 0.165 | 0.18 | 0.175 | 0.17 | 0.2 | 0.932
| Std Dev. |  | 0.005164 | 0.0044 | 0.0063246 | 0.0040825 | 0.008165 | 0.0098319 | 0.0063246 | 0.008165 | 0.0051194 | 0.0032506 | 0.0040825 | 0.0098319 | 0.0063246 | 0.0054772 | 0.0063246 | 0.040825
| Rel.Std.Dev. |  | 2.979% | 2.623% | 3.72% | 2.525% | 4.454% | 5.13% | 3.72% | 4.454% | 2.923% | 2.06% | 2.378% | 5.841% | 3.514% | 3.13% | 3.72% | 22.27%
| PDM3 |  | 0.9393% | -2.303% | -1.002% | -5.855% | 6.763% | 11.62% | -1.002% | 6.763% | 1.999% | -8.087% | -0.0313% | -1.972% | 4.822% | 1.91% | -1.002% | 6.763% | 442.7%
| Z-Score (Absolute) |  | 0.13 | 0.72 | 0.38 | 1.64 | 1.64 | 2.9 | 0.38 | 1.64 | 0.4 | 2.22 | 0.13 | 0.63 | 1.14 | 0.38 | 0.38 | NA | 115.08
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 242b (Certified Value 0.31 ppm) | [21 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.33 | 0.32055 | 0.40851 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.3 | [11 empty] | 1
| 1 | 2 | 0.35 | 0.33654 | 0.36099 | 0.3 | 0.3 | 0.34 | 0.3 | 0.29 | 0.3 | [11 empty] | 4
| 1 | 3 | 0.34 | 0.35123 | 0.39663 | 0.3 | 0.3 | 0.33 | 0.3 | 0.32 | 0.3 | [11 empty] | 16
| 1 | 4 | 0.33 | 0.33596 | 0.40068 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.3 | [11 empty] | 0.31046
| 1 | 5 | 0.33 | 0.32331 | 0.3978 | 0.3 | 0.3 | 0.35 | 0.3 | 0.3 | 0.3 | [11 empty] | 66
| 1 | 6 | 0.33 | 0.32174 | 0.37476 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.3
| Mean |  | 0.335 | 0.33155 | 0.38989 | 0.3 | 0.3 | 0.34 | 0.3 | 0.30667 | 0.3
| Median |  | 0.33 | 0.32963 | 0.39721 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.3
| Std Dev. |  | 0.0083666 | 0.011972 | 0.018085 | 0 | 0 | 0.0063246 | 0 | 0.010328 | 0
| Rel.Std.Dev. |  | 2.497% | 3.611% | 4.638% | 0% | 0% | 1.86% | 0% | 3.368% | 0%
| PDM3 |  | 7.904% | 6.795% | 25.59% | -3.369% | -3.369% | 9.515% | -3.369% | -1.222% | -3.369%
| Z-Score (Absolute) |  | 2.87 | 2.52 | 8.42 | 0.67 | 0.67 | 3.37 | 0.67 | 0 | 0.67
| Analytical results for U in OREAS 242b (Certified Value 0.44 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.42 | 0.4142 | 0.47834 | 0.4 | 0.4 | 0.4 | 0.5 | 0.4 | 0.5 | 0.42408 | 0.41 | 0.45 | 0.4 | 0.47 | 0.4 | 0.44 | 0.5 |  |  |  | 1
| 1 | 2 | 0.44 | 0.42334 | 0.48867 | 0.4 | 0.4 | 0.5 | 0.4 | 0.5 | 0.5 | 0.45355 | 0.4 | 0.46 | 0.4 | 0.46 | 0.4 | 0.42 | 0.6 |  |  |  | 32
| 1 | 3 | 0.44 | 0.40497 | 0.4781 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 | 0.5 | 0.43104 | 0.4 | 0.47 | 0.4 | 0.47 | 0.4 | 0.46 | 0.5 |  |  |  | 16
| 1 | 4 | 0.44 | 0.38651 | 0.49462 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 | 0.4 | 0.45551 | 0.4 | 0.45 | 0.4 | 0.47 | 0.4 | 0.46 | 0.4 |  |  |  | 0.43839
| 1 | 5 | 0.42 | 0.42433 | 0.50078 | 0.4 | 0.4 | 0.4 | 0.5 | 0.4 | 0.4 | 0.43947 | 0.4 | 0.43 | 0.4 | 0.47 | 0.4 | 0.44 | 0.5 |  |  |  | 67
| 1 | 6 | 0.46 | 0.41595 | 0.48423 | 0.4 | 0.4 | 0.5 | 0.4 | 0.4 | 0.5 | 0.4285 | 0.41 | 0.46 | 0.4 | 0.45 | 0.4 | 0.46 | 0.4
| Mean |  | 0.43667 | 0.41155 | 0.48746 | 0.4 | 0.4 | 0.46667 | 0.46667 | 0.45 | 0.46667 | 0.43869 | 0.40333 | 0.45333 | 0.4 | 0.465 | 0.4 | 0.44667 | 0.48333
| Median |  | 0.44 | 0.41508 | 0.48645 | 0.4 | 0.4 | 0.5 | 0.5 | 0.45 | 0.5 | 0.43526 | 0.4 | 0.455 | 0.4 | 0.47 | 0.4 | 0.45 | 0.5
| Std Dev. |  | 0.015055 | 0.014136 | 0.0090705 | 6.0809e-17 | 6.0809e-17 | 0.05164 | 0.05164 | 0.054772 | 0.05164 | 0.013266 | 0.005164 | 0.013663 | 6.0809e-17 | 0.0083666 | 6.0809e-17 | 0.01633 | 0.075277
| Rel.Std.Dev. |  | 3.448% | 3.435% | 1.861% | 1.52e-14% | 1.52e-14% | 11.07% | 11.07% | 12.17% | 11.07% | 3.024% | 1.28% | 3.014% | 1.52e-14% | 1.799% | 1.52e-14% | 3.656% | 15.57%
| PDM3 |  | -0.394% | -6.123% | 11.19% | -8.758% | -8.758% | 6.449% | 6.449% | 2.647% | 6.449% | 0.06777% | -7.998% | 3.408% | -8.758% | 6.069% | -8.758% | 1.887% | 10.25%
| Z-Score (Absolute) |  | 0.34 | 1.18 | 1.38 | 1.57 | 1.57 | 0.67 | 0.67 | 0.11 | 0.67 | 0.27 | 1.46 | 0.22 | 1.57 | 0.62 | 1.57 | 0 | 1.24
| Analytical results for V in OREAS 242b (Certified Value 257 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 260 | 262.4299 | 259.8435 | 149 | 274 | 257 | 252 | 267 | 244 | 247 | 258.5989 | 272 | 264 | 258 | 249 | 251 | 267 | 289.3 | 250 | 259.445 | 1
| 1 | 2 | 270 | 255.9808 | 258.3945 | 124 | 273 | 254 | 256 | 262 | 263 | 249 | 266.0855 | 276 | 258 | 265 | 246 | 262 | 264 | 285.4 | 250 | 260.413 | 33
| 1 | 3 | 263 | 267.6847 | 260.8305 | 118 | 272 | 252 | 255 | 262 | 255 | 249 | 255.4314 | 273 | 258 | 250 | 252 | 255 | 247 | 290.6 | 250 | 91.8 | 16
| 1 | 4 | 264 | 259.6219 | 259.791 | 181 | 274 | 255 | 249 | 257 | 256 | 247 | 265.6162 | 271 | 265 | 268 | 250 | 257 | 263 | 282.1 | 250 | 260.635 | 256.8232
| 1 | 5 | 269 | 255.093 | 261.24 | 267 | 274 | 256 | 253 | 253 | 256 | 243 | 255.3603 | 271 | 265 | 259 | 245 | 256 | 242 | 290.5 | 255 | 258.086 | 68
| 1 | 6 | 269 | 264.7928 | 261.597 | 262 | 272 | 257 | 256 | 263 | 258 | 242 | 253.9145 | 275 | 256 | 249 | 256 | 256 | 248 | 288.6 | 250 | 255.687
| Mean |  | 265.8333 | 260.9338 | 260.2828 | 183.5 | 273.1667 | 255.1667 | 253.5 | 260.6667 | 255.3333 | 246.1667 | 259.1678 | 273 | 261 | 258.1667 | 249.6667 | 256.1667 | 255.1667 | 287.75 | 250.8333 | 231.011
| Median |  | 266.5 | 261.0259 | 260.337 | 165 | 273.5 | 255.5 | 254 | 262 | 256 | 247 | 257.0152 | 272.5 | 261 | 258.5 | 249.5 | 256 | 255.5 | 288.95 | 250 | 258.7655
| Std Dev. |  | 4.0702 | 4.9613 | 1.1785 | 66.5755 | 0.98319 | 1.9408 | 2.7386 | 4.9261 | 6.2503 | 2.9944 | 5.4003 | 2.0976 | 4.0988 | 7.679 | 4.0332 | 3.5449 | 10.6849 | 3.3543 | 2.0412 | 68.2235
| Rel.Std.Dev. |  | 1.531% | 1.901% | 0.4528% | 36.28% | 0.3599% | 0.7606% | 1.08% | 1.89% | 2.448% | 1.216% | 2.084% | 0.7684% | 1.57% | 2.974% | 1.615% | 1.384% | 4.187% | 1.166% | 0.8138% | 29.53%
| PDM3 |  | 3.508% | 1.601% | 1.347% | -28.55% | 6.364% | -0.645% | -1.294% | 1.497% | -0.5801% | -4.149% | 0.9129% | 6.299% | 1.626% | 0.5231% | -2.787% | -0.2556% | -0.645% | 12.04% | -2.332% | -10.05%
| Z-Score (Absolute) |  | 1.54 | 0.67 | 0.55 | 13.07 | 2.84 | 0.35 | 0.65 | 0.62 | 0.33 | 1.95 | 0.36 | 2.81 | 0.68 | 0.18 | 1.33 | 0.18 | 0.35 | 5.43 | 1.12 | 4.64
| Analytical results for W in OREAS 242b (Certified Value 16.1 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 15.2 | 15.6029 | 16.2 | 16.3 | 17.2 | 15.5 | 16.6 | 15.6 | 16.3 | 15.4736 | 16.18 | 15.8 | 15.6 | 16.4 | 15.1 | 16.2 | < 5 |  |  |  | 1
| 1 | 2 | 15.9 | 15.8061 | 16.1 | 13.9 | 17 | 15.8 | 16.6 | 16.8 | 16.6 | 15.8191 | 19.47 | 15.8 | 15.1 | 16.8 | 15.5 | 15.4 | < 5 |  |  |  | 34
| 1 | 3 | 15.6 | 16.5829 | 16.2 | 14.7 | 17 | 16.2 | 16.4 | 17 | 16.8 | 15.66 | 17.7 | 15.6 | 15.3 | 17 | 15.3 | 16.7 | < 5 |  |  |  | 16
| 1 | 4 | 15.5 | 16.022 | 16 | 15.2 | 16.6 | 16.8 | 17 | 16.9 | 16.4 | 16.0083 | 17.78 | 16.1 | 15.6 | 16.5 | 15.1 | 16.8 | < 5 |  |  |  | 16.0697
| 1 | 5 | 15.2 | 15.3797 | 15.9 | 14.3 | 16.6 | 15.2 | 16.4 | 16.8 | 16.5 | 16.0114 | 17.86 | 15.9 | 15.2 | 16.6 | 15.2 | 16.4 | < 5 |  |  |  | 69
| 1 | 6 | 16 | 15.8387 | 15.3 | 14.9 | 16.5 | 16.1 | 16.6 | 15.6 | 16.6 | 15.5175 | 17.47 | 15.9 | 14.8 | 16.6 | 15 | 16.1 | < 5
| Mean |  | 15.5667 | 15.8721 | 15.95 | 14.8833 | 16.8167 | 15.9333 | 16.6 | 16.45 | 16.5333 | 15.7483 | 17.7433 | 15.85 | 15.2667 | 16.65 | 15.2 | 16.2667
| Median |  | 15.55 | 15.8224 | 16.05 | 14.8 | 16.8 | 15.95 | 16.6 | 16.8 | 16.55 | 15.7396 | 17.74 | 15.85 | 15.25 | 16.6 | 15.15 | 16.3
| Std Dev. |  | 0.33862 | 0.41183 | 0.33912 | 0.83046 | 0.28577 | 0.56451 | 0.21909 | 0.66257 | 0.17512 | 0.23596 | 1.0504 | 0.16432 | 0.30768 | 0.21679 | 0.17889 | 0.50465
| Rel.Std.Dev. |  | 2.175% | 2.595% | 2.126% | 5.58% | 1.699% | 3.543% | 1.32% | 4.028% | 1.059% | 1.498% | 5.92% | 1.037% | 2.015% | 1.302% | 1.177% | 3.102%
| PDM3 |  | -3.13% | -1.23% | -0.7448% | -7.383% | 4.648% | -0.8485% | 3.3% | 2.367% | 2.885% | -2% | 10.41% | -1.367% | -4.997% | 3.611% | -5.412% | 1.226%
| Z-Score (Absolute) |  | 0.41 | 0.07 | 0.02 | 1.18 | 0.99 | 0 | 0.75 | 0.58 | 0.67 | 0.21 | 2.03 | 0.09 | 0.75 | 0.81 | 0.82 | 0.37 | 15.1
| Analytical results for Y in OREAS 242b (Certified Value 20.1 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 20.39 | 20.2343 | 22.0934 | 19.2 | 20.9 | 22.2 | 20.7 | 20.1 | 20.5 | 20.1 | 20.031 | 19.74 | 19.6 | 19.7 | 19.33 | 19.5 | 20.8 | 19.3 | 18.884 |  | 1
| 1 | 2 | 20.3 | 19.889 | 20.9252 | 18.8 | 20 | 22.3 | 20.4 | 20.7 | 20.7 | 20.5 | 20.6608 | 19.98 | 20.2 | 19.2 | 19.53 | 20.6 | 20.5 | 19.1 | 18.894 |  | 35
| 1 | 3 | 20.51 | 20.6411 | 21.8315 | 19 | 19.6 | 21.4 | 20.5 | 20.3 | 21.3 | 20.1 | 19.8091 | 19.83 | 20 | 18.9 | 19.96 | 19.7 | 20.5 | 19 | 16.51 |  | 16
| 1 | 4 | 20.52 | 20.2938 | 22.331 | 18.7 | 19.4 | 20.8 | 24.9 | 20 | 21.5 | 20.1 | 20.9087 | 19.58 | 20.5 | 20.3 | 19.82 | 20.1 | 21.4 | 19 | 18.983 |  | 20.067
| 1 | 5 | 20.39 | 19.8733 | 22.2122 | 19.2 | 20.9 | 22 | 21 | 19 | 21.6 | 19.2 | 20.5968 | 19.68 | 19.9 | 19 | 19.8 | 19.8 | 20.3 | 19.7 | 18.686 |  | 70
| 1 | 6 | 20.44 | 20.454 | 21.0215 | 18.6 | 20 | 22.6 | 21 | 20.5 | 20.4 | 20.3 | 19.9079 | 19.96 | 20.3 | 18.5 | 19.66 | 20.3 | 20.8 | 19 | 18.457
| Mean |  | 20.425 | 20.2309 | 21.7358 | 18.9167 | 20.1333 | 21.8833 | 21.4167 | 20.1 | 21 | 20.05 | 20.3191 | 19.795 | 20.0833 | 19.2667 | 19.6833 | 20 | 20.7167 | 19.1833 | 18.4023
| Median |  | 20.415 | 20.2641 | 21.9624 | 18.9 | 20 | 22.1 | 20.85 | 20.2 | 21 | 20.1 | 20.3139 | 19.785 | 20.1 | 19.1 | 19.73 | 19.95 | 20.65 | 19.05 | 18.785
| Std Dev. |  | 0.083126 | 0.30554 | 0.61409 | 0.25626 | 0.6377 | 0.66458 | 1.7244 | 0.59666 | 0.52915 | 0.44609 | 0.45905 | 0.15821 | 0.31885 | 0.64083 | 0.22704 | 0.40988 | 0.38687 | 0.27869 | 0.94608
| Rel.Std.Dev. |  | 0.407% | 1.51% | 2.825% | 1.355% | 3.167% | 3.037% | 8.052% | 2.968% | 2.52% | 2.225% | 2.259% | 0.7992% | 1.588% | 3.326% | 1.153% | 2.049% | 1.867% | 1.453% | 5.141%
| PDM3 |  | 1.784% | 0.8167% | 8.316% | -5.733% | 0.3304% | 9.051% | 6.726% | 0.1643% | 4.649% | -0.08488% | 1.256% | -1.356% | 0.08123% | -3.988% | -1.912% | -0.334% | 3.237% | -4.404% | -8.296%
| Z-Score (Absolute) |  | 0.53 | 0.21 | 2.65 | 1.92 | 0.05 | 2.89 | 2.13 | 0 | 1.46 | 0.08 | 0.35 | 0.49 | 0.03 | 1.35 | 0.67 | 0.16 | 1 | 1.48 | 2.75
| Analytical results for Yb in OREAS 242b (Certified Value 2.16 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.28 | 2.1761 | 2.0641 | 2.3 | 2.2 | 2.16 | 2 | 2.1 | 2.15 | [11 empty] | 1
| 1 | 2 | 2.3 | 2.1557 | 1.9526 | 2.3 | 2.1 | 2.22 | 2.1 | 1.99 | 2.15 | [11 empty] | 5
| 1 | 3 | 2.35 | 2.1721 | 2.0442 | 2.4 | 2.1 | 2.27 | 2 | 2.12 | 2.15 | [11 empty] | 16
| 1 | 4 | 2.23 | 2.0956 | 2.0517 | 2.4 | 2.2 | 2.13 | 2 | 2.13 | 2.15 | [11 empty] | 2.1629
| 1 | 5 | 2.27 | 2.1614 | 2.03 | 2.3 | 2.2 | 2.3 | 2 | 2.1 | 2.25 | [11 empty] | 71
| 1 | 6 | 2.3 | 2.0407 | 1.981 | 2.2 | 2.1 | 2.23 | 2.1 | 2.11 | 2.2
| Mean |  | 2.2883 | 2.1336 | 2.0206 | 2.3167 | 2.15 | 2.2183 | 2.0333 | 2.0917 | 2.175
| Median |  | 2.29 | 2.1586 | 2.0371 | 2.3 | 2.15 | 2.225 | 2 | 2.105 | 2.15
| Std Dev. |  | 0.039707 | 0.054089 | 0.044057 | 0.075277 | 0.054772 | 0.064317 | 0.05164 | 0.051153 | 0.041833
| Rel.Std.Dev. |  | 1.735% | 2.535% | 2.18% | 3.249% | 2.548% | 2.899% | 2.54% | 2.446% | 1.923%
| PDM3 |  | 5.797% | -1.356% | -6.581% | 7.107% | -0.5982% | 2.561% | -5.992% | -3.295% | 0.5576%
| Z-Score (Absolute) |  | 1.37 | 0.16 | 1.28 | 1.64 | 0 | 0.67 | 1.15 | 0.58 | 0.25
| Analytical results for Zn in OREAS 242b (Certified Value 115 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 116 | 114.978 | 119.424 | 106 | 120 | 122 | 117 | 126 | 111 | 113 | 115.7938 | 101 | 117 | 109.4 | 109 | 103.4 | 95.4 | 128.8 | 120 | 109.807 | 1
| 1 | 2 | 114 | 113.5361 | 121.02 | 103 | 120 | 120 | 119 | 123 | 128 | 115 | 119.6786 | 105 | 116 | 113.7 | 108 | 104.7 | 93.2 | 131.1 | 116 | 108.959 | 15
| 1 | 3 | 118 | 118.3768 | 119.94 | 103 | 119 | 119 | 120 | 121 | 117 | 114 | 115.1845 | 103 | 118 | 110.9 | 112 | 104.7 | 95 | 132.1 | 118 | 69.557 | 16
| 1 | 4 | 118 | 114.7027 | 119.004 | 105 | 123 | 120 | 117 | 120 | 117 | 114 | 119.0181 | 103 | 118 | 116.2 | 109 | 107.2 | 97.9 | 131.8 | 118 | 110.652 | 115.3052
| 1 | 5 | 115 | 113.5121 | 119.988 | 106 | 121 | 120 | 121 | 118 | 118 | 113 | 115.8307 | 102 | 117 | 111.1 | 116 | 104.9 | 95.3 | 132.5 | 124 | 109.978 | 72
| 1 | 6 | 118 | 115.9913 | 120.324 | 105 | 123 | 122 | 119 | 123 | 119 | 112 | 115.3729 | 107 | 118 | 106.1 | 109 | 106.5 | 97.6 | 130.9 | 118 | 109.198
| Mean |  | 116.5 | 115.1828 | 119.95 | 104.6667 | 121 | 120.5 | 118.8333 | 121.8333 | 118.3333 | 113.5 | 116.8131 | 103.5 | 117.3333 | 111.2333 | 110.5 | 105.2333 | 95.7333 | 131.2 | 119 | 103.0252
| Median |  | 117 | 114.8403 | 119.964 | 105 | 120.5 | 120 | 119 | 122 | 117.5 | 113.5 | 115.8123 | 103 | 117.5 | 111 | 109 | 104.8 | 95.35 | 131.45 | 118 | 109.5025
| Std Dev. |  | 1.7607 | 1.8238 | 0.70113 | 1.3663 | 1.6733 | 1.2247 | 1.6021 | 2.7869 | 5.5015 | 1.0488 | 1.9902 | 2.1679 | 0.8165 | 3.4812 | 3.0166 | 1.3794 | 1.7569 | 1.3206 | 2.7568 | 16.4069
| Rel.Std.Dev. |  | 1.511% | 1.583% | 0.5845% | 1.305% | 1.383% | 1.016% | 1.348% | 2.287% | 4.649% | 0.9241% | 1.704% | 2.095% | 0.6959% | 3.13% | 2.73% | 1.311% | 1.835% | 1.007% | 2.317% | 15.93%
| PDM3 |  | 1.036% | -0.1061% | 4.028% | -9.226% | 4.939% | 4.505% | 3.06% | 5.662% | 2.626% | -1.566% | 1.308% | -10.24% | 1.759% | -3.531% | -4.167% | -8.735% | -16.97% | 13.79% | 3.204% | -10.65%
| Z-Score (Absolute) |  | 0.03 | 0.24 | 0.54 | 1.98 | 0.72 | 0.63 | 0.36 | 0.85 | 0.28 | 0.52 | 0.03 | 2.17 | 0.11 | 0.89 | 1.01 | 1.88 | 3.45 | 2.4 | 0.39 | 2.25
| Analytical results for Zr in OREAS 242b (Certified Value 55 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
1
| 1 | 1 | 57 | 54.9944 | 62.525 | 30 | 56 | 58.1 | 55 | 59.5 | 55.4 | 50.6 | 53.6817 | 50.5 | 55.2 | 55.4 | 54 | 48.5 | 56 | 39.06 | 35 |  | 1
| 1 | 2 | 53.3 | 54.41 | 60.4985 | 15 | 55 | 62.1 | 51.9 | 52 | 60.1 | 49.9 | 53.8772 | 51.2 | 55.9 | 49.1 | 51.9 | 51.7 | 54 | 40.49 | 40 |  | 16
| 1 | 3 | 54.9 | 56.0844 | 60.11 | 14 | 57 | 60.8 | 51.1 | 57.4 | 61.4 | 50.8 | 52.7139 | 50.5 | 55.8 | 48.9 | 52.5 | 53.9 | 58 | 42.12 | 40 |  | 16
| 1 | 4 | 57 | 56.2212 | 61.496 | 40 | 57 | 56.6 | 51.4 | 52.3 | 61.4 | 49 | 54.3224 | 49.7 | 54.3 | 51.7 | 55.4 | 49.9 | 57 | 40.75 | 40 |  | 54.6677
| 1 | 5 | 51.9 | 53.4393 | 62.084 | 54 | 54 | 64.1 | 57.3 | 50.2 | 59.1 | 51.5 | 52.9416 | 50.6 | 53.8 | 50.2 | 53.5 | 49 | 57 | 40.99 | 40 |  | 73
| 1 | 6 | 55.9 | 56.8349 | 60.7715 | 51 | 56 | 56 | 53 | 54.6 | 67.9 | 55.2 | 54.3545 | 50.5 | 54.4 | 46.7 | 52 | 63.6 | 57 | 40.66 | 40
| Mean |  | 55 | 55.3307 | 61.2475 | 34 | 55.8333 | 59.6167 | 53.2833 | 54.3333 | 60.8833 | 51.1667 | 53.6486 | 50.5 | 54.9 | 50.3333 | 53.2167 | 52.7667 | 56.5 | 40.6783 | 39.1667
| Median |  | 55.4 | 55.5394 | 61.1338 | 35 | 56 | 59.45 | 52.45 | 53.45 | 60.75 | 50.7 | 53.7794 | 50.5 | 54.8 | 49.65 | 53 | 50.8 | 57 | 40.705 | 40
| Std Dev. |  | 2.0649 | 1.2776 | 0.94604 | 17.3321 | 1.169 | 3.2295 | 2.4244 | 3.5449 | 4.0897 | 2.151 | 0.68981 | 0.47749 | 0.86255 | 2.9777 | 1.3556 | 5.6638 | 1.3784 | 0.98341 | 2.0412
| Rel.Std.Dev. |  | 3.754% | 2.309% | 1.545% | 50.98% | 2.094% | 5.417% | 4.55% | 6.524% | 6.717% | 4.204% | 1.286% | 0.9455% | 1.571% | 5.916% | 2.547% | 10.73% | 2.44% | 2.418% | 5.212%
| PDM3 |  | 0.6078% | 1.213% | 12.04% | -37.81% | 2.132% | 9.053% | -2.532% | -0.6117% | 11.37% | -6.404% | -1.864% | -7.624% | 0.4249% | -7.929% | -2.654% | -3.477% | 3.352% | -25.59% | -28.36%
| Z-Score (Absolute) |  | 0.37 | 0.46 | 2.06 | 5.34 | 0.59 | 1.62 | 0.1 | 0.19 | 1.97 | 0.67 | 0 | 0.86 | 0.34 | 0.9 | 0.12 | 0.24 | 0.77 | 3.52 | 3.93
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 242b (Certified Value 4.58 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 4.3 | 4.5679 | 4.4209 | 4.82 | 4.9 | 4.73 | 4.61 | 4.46 | 4.72 | 4.74 | 4.5856 | 4.3 | 4.59 | 3.803 | 4.5 | 4.522 | 4.46 | 6.7 | 5.15 | 3.963 |  | 1
| 1 | 2 | 4.52 | 4.4918 | 4.1045 | 4.71 | 4.46 | 4.6 | 4.68 | 4.53 | 4.73 | 4.58 | 4.738 | 4.3 | 4.57 | 3.734 | 4.6 | 4.495 | 4.64 | 6.61 | 0.35 | 4.153 |  | 22
| 1 | 3 | 4.4 | 4.6227 | 4.338 | 4.76 | 4.5 | 4.55 | 4.6 | 4.48 | 4.75 | 4.66 | 4.7241 | 4.4 | 4.51 | 3.937 | 4.6 | 4.568 | 4.59 | 6.73 | 4.55 | 4.276 |  | 16
| 1 | 4 | 4.48 | 4.6046 | 4.2464 | 4.78 | 4.73 | 4.68 | 4.69 | 4.41 | 4.78 | 4.64 | 4.6776 | 4.3 | 4.4 | 3.825 | 4.6 | 4.398 | 4.55 | 6.61 | 5.1 | 4.101 |  | 4.5837
| 1 | 5 | 4.36 | 4.591 | 4.5024 | 4.79 | 4.65 | 4.47 | 4.62 | 4.5 | 4.78 | 4.68 | 4.598 | 4.3 | 4.49 | 3.78 | 4.8 | 4.526 | 4.61 | 6.92 | 4.7 | 4.03 |  | 75
| 1 | 6 | 4.44 | 4.5903 | 4.7015 | 4.65 | 4.44 | 4.52 | 4.68 | 4.53 | 4.84 | 4.71 | 4.7734 | 4.4 | 4.58 | 4.031 | 4.7 | 4.51 | 4.51 | 6.42 | 5.2 | 4.024
| Mean |  | 4.4167 | 4.5781 | 4.3856 | 4.7517 | 4.6133 | 4.5917 | 4.6467 | 4.485 | 4.7667 | 4.6683 | 4.6828 | 4.3333 | 4.5233 | 3.8517 | 4.6333 | 4.5032 | 4.56 | 6.665 | 4.175 | 4.0912
| Median |  | 4.42 | 4.5907 | 4.3794 | 4.77 | 4.575 | 4.575 | 4.65 | 4.49 | 4.765 | 4.67 | 4.7008 | 4.3 | 4.54 | 3.814 | 4.6 | 4.516 | 4.57 | 6.655 | 4.9 | 4.0655
| Std Dev. |  | 0.080416 | 0.045954 | 0.2075 | 0.061779 | 0.18085 | 0.098675 | 0.040825 | 0.045935 | 0.043665 | 0.056006 | 0.076962 | 0.05164 | 0.072572 | 0.11091 | 0.10328 | 0.057014 | 0.066933 | 0.16526 | 1.8923 | 0.11204
| Rel.Std.Dev. |  | 1.821% | 1.004% | 4.731% | 1.3% | 3.92% | 2.149% | 0.8786% | 1.024% | 0.9161% | 1.2% | 1.644% | 1.192% | 1.604% | 2.879% | 2.229% | 1.266% | 1.468% | 2.479% | 45.32% | 2.739%
| PDM3 |  | -3.644% | -0.1231% | -4.322% | 3.664% | 0.6464% | 0.1737% | 1.374% | -2.153% | 3.992% | 1.846% | 2.162% | -5.462% | -1.317% | -15.97% | 1.083% | -1.757% | -0.5171% | 45.41% | -8.916% | -10.75%
| Z-Score (Absolute) |  | 0.96 | 0.06 | 1.16 | 1.16 | 0.28 | 0.14 | 0.49 | 0.53 | 1.25 | 0.63 | 0.72 | 1.49 | 0.29 | 4.54 | 0.41 | 0.42 | 0.06 | 13.27 | 2.49 | 3.02
| Analytical results for Al in OREAS 242b (Certified Value 3 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 2.9178 | 2.958 | 2.94 | 2.99 | 2.96 | 3.1 | 2.93 | 2.98 | 3.08 | 3.2078 | 3.85 | 3.09 | 3.1432 | 2.76 | 3.07 | 2.81 | 3.02 | 3.0043 |  |  |  | 1
| 1 | 2 | 3.0697 | 3.0492 | 2.84 | 2.97 | 2.94 | 3.1 | 3 | 3.01 | 3.08 | 3.0832 | 3.79 | 2.83 | 3.2243 | 2.8 | 2.95 | 2.89 | 2.98 | 3.0175 |  |  |  | #N/A
| 1 | 3 | 3.0167 | 3.1025 | 2.89 | 2.99 | 2.97 | 3.1 | 2.94 | 3.01 | 3.08 | 3.0832 | 3.73 | 2.96 | 3.1843 | 2.83 | 2.91 | 2.85 | 2.98 | 3.0418 |  |  |  | 16
| 1 | 4 | 3.0686 | 3.0331 | 2.9 | 2.98 | 3.04 | 3.13 | 2.93 | 2.99 | 3.09 | 3.1527 | 3.72 | 2.79 | 3.156 | 2.68 | 2.74 | 2.88 | 3.01 | 3.1162 |  |  |  | 2.9964
| 1 | 5 | 3.064 | 3.0647 | 2.92 | 2.99 | 2.97 | 3.11 | 2.98 | 3.03 | 3.06 | 3.2003 | 3.77 | 2.96 | 3.2719 | 2.85 | 2.7 | 2.86 | 2.98 | 2.9169 |  |  |  | 76
| 1 | 6 | 3.0905 | 3.0794 | 2.88 | 2.98 | 2.97 | 3.09 | 2.98 | 2.97 | 3.14 | 3.0888 | 3.79 | 2.94 | 3.2454 | 2.83 | 2.98 | 2.84 | 3.03 | 3.0032
| Mean |  | 3.0379 | 3.0478 | 2.895 | 2.9833 | 2.975 | 3.105 | 2.96 | 2.9983 | 3.0883 | 3.136 | 3.775 | 2.9283 | 3.2042 | 2.7917 | 2.8917 | 2.855 | 3 | 3.0167
| Median |  | 3.0663 | 3.057 | 2.895 | 2.985 | 2.97 | 3.1 | 2.96 | 3 | 3.08 | 3.1208 | 3.78 | 2.95 | 3.2043 | 2.815 | 2.93 | 2.855 | 2.995 | 3.0109
| Std Dev. |  | 0.063679 | 0.050101 | 0.034496 | 0.008165 | 0.033912 | 0.013784 | 0.030332 | 0.022286 | 0.027142 | 0.058931 | 0.047223 | 0.10685 | 0.051235 | 0.063061 | 0.14359 | 0.02881 | 0.022804 | 0.064569
| Rel.Std.Dev. |  | 2.096% | 1.644% | 1.192% | 0.2737% | 1.14% | 0.4439% | 1.025% | 0.7433% | 0.8788% | 1.879% | 1.251% | 3.649% | 1.599% | 2.259% | 4.965% | 1.009% | 0.7601% | 2.14%
| PDM3 |  | 1.386% | 1.717% | -3.383% | -0.435% | -0.7131% | 3.626% | -1.214% | 0.06564% | 3.069% | 4.66% | 25.99% | -2.271% | 6.936% | -6.832% | -3.494% | -4.718% | 0.1213% | 0.6772%
| Z-Score (Absolute) |  | 0.33 | 0.41 | 0.88 | 0.13 | 0.2 | 0.89 | 0.33 | 0.01 | 0.75 | 1.15 | 6.54 | 0.6 | 1.73 | 1.75 | 0.91 | 1.22 | 0.01 | 0.15
| Analytical results for As in OREAS 242b (Certified Value 51 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 1
| 1 | 1 | 49 | 50.6825 | 50.75 | 47.646 | 52.5 | 55 | 53.6 | 54.2 | 50.1 | 52.7 | 51.7 | 51.4238 | 48.9 | 51 | 50.7 | 52 | 49.7 | 51.3 | 41.7 | 46 | 49.335 | 1
| 1 | 2 | 52 | 49.7473 | 52.538 | 47.833 | 51.3 | 52 | 51.9 | 55.9 | 51 | 53.6 | 51.9 | 49.2802 | 48.4 | 51 | 47.5 | 53 | 50 | 51.8 | 42 | 47 | 49.286 | 10
| 1 | 3 | 51 | 51.7265 | 48.28 | 48.13 | 51.5 | 53 | 50.8 | 55.2 | 50.6 | 53.8 | 51.6 | 49.5137 | 48.4 | 52 | 48.7 | 53 | 50.2 | 51.3 | 41.3 | 47 | 49.569 | 16
| 1 | 4 | 52 | 50.7195 | 47.462 | 49.549 | 52 | 51 | 53.4 | 56.3 | 49.8 | 54.2 | 52.1 | 49.7654 | 48.5 | 51 | 48.2 | 52 | 49 | 50.9 | 39.3 | 46.5 | 49.264 | 50.8458
| 1 | 5 | 50 | 50.5473 | 52.178 | 48.163 | 52 | 53 | 50.2 | 54.7 | 50.8 | 53.9 | 52.7 | 51.6052 | 48.3 | 51 | 49.8 | 54 | 50.3 | 52 | 41.5 | 45.5 | 48.584 | 77
| 1 | 6 | 52 | 50.7975 | 48.235 | 48.218 | 50.1 | 52 | 51.1 | 55.7 | 51.8 | 54.5 | 53.4 | 49.8093 | 49.1 | 52 | 48 | 54 | 49.7 | 51.3 | 41.9 | 47.5 | 48.757
| Mean |  | 51 | 50.7034 | 49.9072 | 48.2565 | 51.5667 | 52.6667 | 51.8333 | 55.3333 | 50.6833 | 53.7833 | 52.2333 | 50.2329 | 48.6 | 51.3333 | 48.8167 | 53 | 49.8167 | 51.4333 | 41.2833 | 46.5833 | 49.1325
| Median |  | 51.5 | 50.701 | 49.515 | 48.1465 | 51.75 | 52.5 | 51.5 | 55.45 | 50.7 | 53.85 | 52 | 49.7873 | 48.45 | 51 | 48.45 | 53 | 49.85 | 51.3 | 41.6 | 46.75 | 49.275
| Std Dev. |  | 1.2649 | 0.63163 | 2.2005 | 0.67065 | 0.83347 | 1.3663 | 1.4038 | 0.78655 | 0.70545 | 0.61779 | 0.69186 | 1.0123 | 0.32249 | 0.5164 | 1.209 | 0.89443 | 0.47081 | 0.39833 | 1.0048 | 0.73598 | 0.37795
| Rel.Std.Dev. |  | 2.48% | 1.246% | 4.409% | 1.39% | 1.616% | 2.594% | 2.708% | 1.421% | 1.392% | 1.149% | 1.325% | 2.015% | 0.6636% | 1.006% | 2.477% | 1.688% | 0.9451% | 0.7745% | 2.434% | 1.58% | 0.7693%
| PDM3 |  | 0.3033% | -0.28% | -1.846% | -5.092% | 1.418% | 3.581% | 1.942% | 8.826% | -0.3195% | 5.777% | 2.729% | -1.205% | -4.417% | 0.9589% | -3.991% | 4.237% | -2.024% | 1.156% | -18.81% | -8.383% | -3.37%
| Z-Score (Absolute) |  | 0.13 | 0 | 0.35 | 1.08 | 0.38 | 0.87 | 0.5 | 2.04 | 0.01 | 1.36 | 0.67 | 0.21 | 0.93 | 0.28 | 0.83 | 1.01 | 0.39 | 0.32 | 4.15 | 1.82 | 0.69
| Analytical results for B in OREAS 242b (Certified Value 14.7 ppm) | [22 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | [8 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [8 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 15g | [8 empty] | 1
| 1 | 1 |  | 15.0114 | 20 | 15 | 10 | 20 | 10 | 10 | 10 | 17.4774 | < 20 | 24 | 12 | [8 empty] | 1
| 1 | 2 | 13 | 15.5999 | 19 | 17 | 10 | 20 | 10 | 10 | 10 | 17.4172 | < 20 | 23 | 12 | [8 empty] | 1
| 1 | 3 | 13 | 15.4739 | 19 | 17 | 10 | 20 | 10 | 10 | 10 | 16.1305 | < 20 | 24 | 12 | [8 empty] | 16
| 1 | 4 | 13 | 14.766 | 19 | 15 | 10 | 20 | 10 | 10 | 10 | 18.1567 | < 20 | 23 | 11 | [8 empty] | 14.7145
| 1 | 5 | 13 | 14.5365 | 19 | 17 | 10 | 20 | 10 | 10 | 10 | 18.8884 | < 20 | 25 | 12 | [8 empty] | 78
| 1 | 6 | 13 | 15.9398 | 19 | 16 | 10 | 20 | 10 | 10 | 10 | 16.0475 | < 20 | 24 | 12
| Mean |  | 13 | 15.2212 | 19.1667 | 16.1667 | 10 | 20 | 10 | 10 | 10 | 17.3529 |  | 23.8333 | 11.8333
| Median |  | 13 | 15.2426 | 19 | 16.5 | 10 | 20 | 10 | 10 | 10 | 17.4473 |  | 24 | 12
| Std Dev. |  | 0 | 0.53734 | 0.40825 | 0.98319 | 0 | 0 | 0 | 0 | 0 | 1.1155 |  | 0.75277 | 0.40825
| Rel.Std.Dev. |  | 0% | 3.53% | 2.13% | 6.082% | 0% | 0% | 0% | 0% | 0% | 6.428% |  | 3.158% | 3.45%
| PDM3 |  | -11.65% | 3.444% | 30.26% | 9.869% | -32.04% | 35.92% | -32.04% | -32.04% | -32.04% | 17.93% |  | 61.97% | -19.58%
| Z-Score (Absolute) |  | 0 | 0.5 | 1.39 | 0.71 | 0.67 | 1.57 | 0.67 | 0.67 | 0.67 | 0.98 | 0.67 | 2.44 | 0.26
| Analytical results for Ba in OREAS 242b (Certified Value 30.3 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 0.5g |  |  |  | 2
| 1 | 1 | 24 | 29.3853 | 31.517 | 35.8 | 32 | 30 | 30 | 30 | 30 | 30 | 29.4742 | 36 | 31.2 | 37 | 26.1 | 33.3 | 27.3 | 36.875 |  |  |  | 1
| 1 | 2 | 30 | 29.3595 | 33.272 | 34.5 | 31 | 30 | 30 | 30 | 30 | 30 | 28.348 | 36 | 29 | 37 | 25.8 | 35.2 | 28.9 | 34.807 |  |  |  | 24
| 1 | 3 | 29 | 30.4605 | 33.004 | 35.8 | 32 | 30 | 30 | 30 | 30 | 30 | 28.3143 | 37 | 30.1 | 36 | 25.3 | 31.9 | 28.3 | 34.533 |  |  |  | 16
| 1 | 4 | 29 | 29.8025 | 33.13 | 35.2 | 33 | 30 | 30 | 30 | 30 | 30 | 29.6892 | 37 | 29.1 | 36 | 24.4 | 30.9 | 27.9 | 35.015 |  |  |  | 30.2786
| 1 | 5 | 29 | 29.8621 | 33.414 | 34.1 | 33 | 30 | 30 | 30 | 30 | 30 | 29.0182 | 36 | 30.5 | 38 | 25.9 | 29.2 | 29.9 | 34.158 |  |  |  | 79
| 1 | 6 | 30 | 29.6168 | 33.442 | 36.2 | 32 | 30 | 30 | 30 | 30 | 30 | 28.1431 | 36 | 30.2 | 37 | 24.9 | 32.3 | 28.2 | 33.498
| Mean |  | 28.5 | 29.7478 | 32.9632 | 35.2667 | 32.1667 | 30 | 30 | 30 | 30 | 30 | 28.8312 | 36.3333 | 30.0167 | 36.8333 | 25.4 | 32.1333 | 28.4167 | 34.8143
| Median |  | 29 | 29.7097 | 33.201 | 35.5 | 32 | 30 | 30 | 30 | 30 | 30 | 28.6831 | 36 | 30.15 | 37 | 25.55 | 32.1 | 28.25 | 34.67
| Std Dev. |  | 2.2583 | 0.40579 | 0.72785 | 0.82381 | 0.75277 | 0 | 0 | 0 | 0 | 0 | 0.65709 | 0.5164 | 0.84242 | 0.75277 | 0.65727 | 2.0481 | 0.89536 | 1.1427
| Rel.Std.Dev. |  | 7.924% | 1.364% | 2.208% | 2.336% | 2.34% | 0% | 0% | 0% | 0% | 0% | 2.279% | 1.421% | 2.806% | 2.044% | 2.588% | 6.374% | 3.151% | 3.282%
| PDM3 |  | -5.874% | -1.753% | 8.866% | 16.47% | 6.236% | -0.9201% | -0.9201% | -0.9201% | -0.9201% | -0.9201% | -4.78% | 20% | -0.8651% | 21.65% | -16.11% | 6.126% | -6.149% | 14.98%
| Z-Score (Absolute) |  | 0.66 | 0.11 | 1.3 | 2.3 | 0.95 | 0 | 0 | 0 | 0 | 0 | 0.51 | 2.77 | 0.01 | 2.99 | 2.01 | 0.93 | 0.69 | 2.11
| Analytical results for Be in OREAS 242b (Certified Value 0.2 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.17 | < 0.5 | 0.772 | 0.2 | 0.21 | 0.2 | 0.21 | 0.18 | 0.2 | 0.18 | < 0.5 | < 0.5 | 0.27 | < 2 | 0.018 | [6 empty] | 1
| 1 | 2 | 0.19 | < 0.5 | 0.86 | 0.2 | 0.2 | 0.21 | 0.22 | 0.17 | 0.21 | 0.18 | < 0.5 | < 0.5 | 0.25 | < 2 | 0.017 | [6 empty] | 25
| 1 | 3 | 0.17 | < 0.5 | 0.873 | 0.2 | 0.21 | 0.2 | 0.21 | 0.17 | 0.21 | 0.18 | < 0.5 | < 0.5 | 0.26 | < 2 | 0.018 | [6 empty] | 16
| 1 | 4 | 0.21 | < 0.5 | 0.84 | 0.2 | 0.23 | 0.21 | 0.21 | 0.17 | 0.21 | 0.18 | < 0.5 | < 0.5 | 0.26 | < 2 | [7 empty] | 0.20375
| 1 | 5 | 0.2 | < 0.5 | 0.875 | 0.2 | 0.18 | 0.2 | 0.21 | 0.17 | 0.21 | 0.18 | < 0.5 | < 0.5 | 0.24 | < 2 | 0.025 | [6 empty] | 80
| 1 | 6 | 0.21 | < 0.5 | 0.836 | 0.2 | 0.2 | 0.2 | 0.22 | 0.17 | 0.21 | 0.18 | < 0.5 | < 0.5 | 0.26 | < 2 | 0.031
| Mean |  | 0.19167 |  | 0.84267 | 0.2 | 0.205 | 0.20333 | 0.21333 | 0.17167 | 0.20833 | 0.18 |  |  | 0.25667 |  | 0.0218
| Median |  | 0.195 |  | 0.85 | 0.2 | 0.205 | 0.2 | 0.21 | 0.17 | 0.21 | 0.18 |  |  | 0.26 |  | 0.018
| Std Dev. |  | 0.018348 |  | 0.03824 | 3.0405e-17 | 0.016432 | 0.005164 | 0.005164 | 0.0040825 | 0.0040825 | 3.0405e-17 |  |  | 0.010328 |  | 0.0060581
| Rel.Std.Dev. |  | 9.573% |  | 4.538% | 1.52e-14% | 8.015% | 2.54% | 2.421% | 2.378% | 1.96% | 1.689e-14% |  |  | 4.024% |  | 27.79%
| PDM3 |  | -5.93% |  | 313.6% | -1.84% | 0.6135% | -0.2045% | 4.703% | -15.75% | 2.249% | -11.66% |  |  | 25.97% |  | -89.3%
| Z-Score (Absolute) |  | 0.33 | 0.67 | 10.88 | NA | 0.1 | 0.13 | 0.04 | 0.67 | 0.04 | 0.53 | 0.67 | 0.67 | 0.79 | 13.59 | 3.25
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 242b (Certified Value 0.4 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 3
| 1 | 1 | 0.39 | < 3.5 | 0.40563 | 0.43 | 0.52 | 0.39 | 0.43 | 0.38 | 0.38 | 0.38 | 0.39814 | < 0.02 | 0.42 | 0.38 | 0.41 | 0.43 | < 5 | 0.3 |  |  |  | 1
| 1 | 2 | 0.4 | < 3.5 | 0.3928 | 0.43 | 0.54 | 0.38 | 0.42 | 0.38 | 0.38 | 0.39 | 0.37647 | < 0.02 | 0.41 | 0.34 | 0.39 | 0.43 | < 5 | 0.4 |  |  |  | 26
| 1 | 3 | 0.4 | < 3.5 | 0.4072 | 0.41 | 0.51 | 0.38 | 0.42 | 0.39 | 0.39 | 0.42 | 0.37614 | < 0.02 | 0.42 | 0.39 | 0.39 | 0.43 | < 5 | 0.4 |  |  |  | 16
| 1 | 4 | 0.41 | < 3.5 | 0.40617 | 0.42 | 0.52 | 0.39 | 0.42 | 0.38 | 0.38 | 0.41 | 0.387 | < 0.02 | 0.42 | 0.37 | 0.38 | 0.45 | < 5 | 0.3 |  |  |  | 0.39971
| 1 | 5 | 0.39 | < 3.5 | 0.41235 | 0.41 | 0.52 | 0.37 | 0.42 | 0.39 | 0.39 | 0.4 | 0.3919 | < 0.02 | 0.42 | 0.38 | 0.4 | 0.44 | < 5 | 0.3 |  |  |  | 81
| 1 | 6 | 0.41 | < 3.5 | 0.3926 | 0.4 | 0.5 | 0.38 | 0.41 | 0.39 | 0.39 | 0.38 | 0.391 | < 0.02 | 0.41 | 0.38 | 0.38 | 0.42 | < 5 | 0.3
| Mean |  | 0.4 |  | 0.40279 | 0.41667 | 0.51833 | 0.38167 | 0.42 | 0.385 | 0.385 | 0.39667 | 0.38677 |  | 0.41667 | 0.37333 | 0.39167 | 0.43333 |  | 0.33333
| Median |  | 0.4 |  | 0.4059 | 0.415 | 0.52 | 0.38 | 0.42 | 0.385 | 0.385 | 0.395 | 0.389 |  | 0.42 | 0.38 | 0.39 | 0.43 |  | 0.3
| Std Dev. |  | 0.0089443 |  | 0.0081722 | 0.012111 | 0.013292 | 0.0075277 | 0.0063246 | 0.0054772 | 0.0054772 | 0.01633 | 0.0088618 |  | 0.005164 | 0.017512 | 0.01169 | 0.010328 |  | 0.05164
| Rel.Std.Dev. |  | 2.236% |  | 2.029% | 2.907% | 2.564% | 1.972% | 1.506% | 1.423% | 1.423% | 4.117% | 2.291% |  | 1.239% | 4.691% | 2.985% | 2.383% |  | 15.49%
| PDM3 |  | 0.0725% |  | 0.7711% | 4.242% | 29.68% | -4.514% | 5.076% | -3.68% | -3.68% | -0.7614% | -3.236% |  | 4.242% | -6.599% | -2.012% | 8.412% |  | -16.61%
| Z-Score (Absolute) |  | 0 | 54.62 | 0.11 | 0.67 | 4.79 | 0.74 | 0.81 | 0.61 | 0.61 | 0.13 | 0.54 | 15.78 | 0.67 | 1.08 | 0.34 | 1.35 | 84.96 | NA
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 242b (Certified Value 2.22 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 2.2 | 2.2462 | 1.4784 | 3.28 | 2.37 | 2.19 | 2.48 | 2.11 | 2.12 | 2.13 | 2.5607 | 1.6869 | 3.22 | 1.95 | 3.2 | 1.87 | 3.28 | 2.11 | 1.51 | 2.238 |  | 1
| 1 | 2 | 2.32 | 2.3535 | 1.5441 | 3.16 | 2.38 | 2.18 | 2.45 | 2.16 | 2.14 | 2.18 | 2.4593 | 1.6962 | 3.18 | 1.77 | 3.33 | 1.94 | 3.29 | 2.26 | 1.46 | 2.2711 |  | #N/A
| 1 | 3 | 2.27 | 2.3661 | 1.5789 | 3.19 | 2.4 | 2.22 | 2.45 | 2.13 | 2.19 | 2.17 | 2.4242 | 1.7066 | 3.26 | 1.84 | 3.25 | 1.92 | 3.25 | 2.18 | 1.46 | 2.3092 |  | 16
| 1 | 4 | 2.32 | 2.2856 | 1.5299 | 3.17 | 2.38 | 2.27 | 2.47 | 2.14 | 2.14 | 2.19 | 2.6608 | 1.6818 | 3.11 | 1.75 | 3.26 | 1.8 | 3.27 | 2.22 | 1.48 | 2.4194 |  | 2.2208
| 1 | 5 | 2.33 | 2.3479 | 1.5817 | 3.16 | 2.36 | 2.21 | 2.45 | 2.16 | 2.17 | 2.16 | 2.6674 | 1.7095 | 3.16 | 1.91 | 3.35 | 1.97 | 3.16 | 2.22 | 1.49 | 2.1266 |  | 82
| 1 | 6 | 2.34 | 2.3521 | 1.5471 | 3.13 | 2.35 | 2.2 | 2.44 | 2.15 | 2.13 | 2.24 | 2.4408 | 1.715 | 3.15 | 1.83 | 3.33 | 1.95 | 3.26 | 2.17 | 1.52 | 2.1821
| Mean |  | 2.2967 | 2.3252 | 1.5434 | 3.1817 | 2.3733 | 2.2117 | 2.4567 | 2.1417 | 2.1483 | 2.1783 | 2.5355 | 1.6993 | 3.18 | 1.8417 | 3.2867 | 1.9083 | 3.2517 | 2.1933 | 1.4867 | 2.2577
| Median |  | 2.32 | 2.35 | 1.5456 | 3.165 | 2.375 | 2.205 | 2.45 | 2.145 | 2.14 | 2.175 | 2.51 | 1.7014 | 3.17 | 1.835 | 3.295 | 1.93 | 3.265 | 2.2 | 1.485 | 2.2545
| Std Dev. |  | 0.053166 | 0.047991 | 0.037813 | 0.051929 | 0.017512 | 0.031885 | 0.015055 | 0.019408 | 0.026394 | 0.03656 | 0.11038 | 0.013217 | 0.053292 | 0.077567 | 0.058878 | 0.063061 | 0.047081 | 0.052026 | 0.025033 | 0.10222
| Rel.Std.Dev. |  | 2.315% | 2.064% | 2.45% | 1.632% | 0.7379% | 1.442% | 0.6128% | 0.9062% | 1.229% | 1.678% | 4.353% | 0.7778% | 1.676% | 4.212% | 1.791% | 3.304% | 1.448% | 2.372% | 1.684% | 4.527%
| PDM3 |  | 3.415% | 4.702% | -30.51% | 43.26% | 6.867% | -0.4127% | 10.62% | -3.565% | -3.265% | -1.914% | 14.17% | -23.48% | 43.19% | -17.07% | 47.99% | -14.07% | 46.42% | -1.238% | -33.06% | 1.661%
| Z-Score (Absolute) |  | 0.16 | 0.23 | 1.78 | 2.44 | 0.36 | 0.06 | 0.57 | 0.24 | 0.22 | 0.15 | 0.78 | 1.38 | 2.44 | 1.01 | 2.71 | 0.84 | 2.62 | 0.11 | 1.93 | 0.06
| Analytical results for Cd in OREAS 242b (Certified Value 0.37 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 0.35 | 0.36818 | 0.33456 | 0.43 | 0.41 | 0.37 | 0.4 | 0.35 | 0.38 | 0.38 | 0.35826 | 0.39 | 0.41 | 0.36 | 0.19 | < 1 | 0.4 | 0.449 |  |  |  | 1
| 1 | 2 | 0.37 | 0.37346 | 0.37277 | 0.37 | 0.41 | 0.38 | 0.39 | 0.34 | 0.41 | 0.38 | 0.35541 | 0.45 | 0.33 | 0.34 | 0.2 | < 1 | 0.4 | 0.297 |  |  |  | 27
| 1 | 3 | 0.33 | 0.35515 | 0.34956 | 0.4 | 0.41 | 0.35 | 0.39 | 0.34 | 0.38 | 0.37 | 0.3633 | 0.4 | 0.42 | 0.35 | 0.2 | < 1 | 0.4 | 0.357 |  |  |  | 16
| 1 | 4 | 0.35 | 0.37287 | 0.3704 | 0.41 | 0.37 | 0.37 | 0.41 | 0.35 | 0.39 | 0.37 | 0.34874 | 0.4 | 0.35 | 0.33 | 0.19 | < 1 | 0.4 | 0.487 |  |  |  | 0.37325
| 1 | 5 | 0.35 | 0.37108 | 0.35562 | 0.36 | 0.41 | 0.37 | 0.38 | 0.36 | 0.39 | 0.36 | 0.38017 | 0.43 | 0.25 | 0.35 | 0.19 | < 1 | 0.4 | 0.349 |  |  |  | 83
| 1 | 6 | 0.36 | 0.34757 | 0.36249 | 0.37 | 0.41 | 0.35 | 0.38 | 0.35 | 0.4 | 0.37 | 0.32657 | 0.39 | 0.39 | 0.35 | 0.21 | < 1 | 0.4 | 0.37
| Mean |  | 0.35167 | 0.36472 | 0.35756 | 0.39 | 0.40333 | 0.365 | 0.39167 | 0.34833 | 0.39167 | 0.37167 | 0.35541 | 0.41 | 0.35833 | 0.34667 | 0.19667 |  | 0.4 | 0.38483
| Median |  | 0.35 | 0.36963 | 0.35905 | 0.385 | 0.41 | 0.37 | 0.39 | 0.35 | 0.39 | 0.37 | 0.35683 | 0.4 | 0.37 | 0.35 | 0.195 |  | 0.4 | 0.3635
| Std Dev. |  | 0.013292 | 0.010778 | 0.014265 | 0.027568 | 0.01633 | 0.012247 | 0.01169 | 0.0075277 | 0.01169 | 0.0075277 | 0.017661 | 0.024495 | 0.063377 | 0.010328 | 0.008165 |  | 6.0809e-17 | 0.07007
| Rel.Std.Dev. |  | 3.78% | 2.955% | 3.989% | 7.069% | 4.049% | 3.355% | 2.985% | 2.161% | 2.985% | 2.025% | 4.969% | 5.974% | 17.69% | 2.979% | 4.152% |  | 1.52e-14% | 18.21%
| PDM3 |  | -5.783% | -2.285% | -4.203% | 4.487% | 8.06% | -2.21% | 4.934% | -6.676% | 4.934% | -0.4244% | -4.78% | 9.846% | -3.997% | -7.122% | -47.31% |  | 7.167% | 3.103%
| Z-Score (Absolute) |  | 0.49 | 0.01 | 0.27 | 0.92 | 1.41 | 0 | 0.98 | 0.61 | 0.98 | 0.25 | 0.35 | 1.66 | 0.25 | 0.67 | 6.19 | 4.97 | NA | 0.73
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 242b (Certified Value 10.8 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 1
| 1 | 1 | 9.514 | 9.559 | 13.5844 | 11.2 | 10.79 | 10.65 | 10.65 | 9.87 | 9.94 | 10.2 | 9.9935 | 12.1 | 15.6 | 14 | 21.471 | [6 empty] | 1
| 1 | 2 | 9.803 | 9.382 | 13.5436 | 10.7 | 10.45 | 10.75 | 10.65 | 9.95 | 9.8 | 10.05 | 9.4186 | 12.4 | 14.9 | 13.6 | 21.482 | [6 empty] | 28
| 1 | 3 | 9.375 | 9.6903 | 13.4572 | 10.8 | 10.5 | 10.25 | 10.7 | 10.35 | 10 | 10.3 | 9.3551 | 12.2 | 15.8 | 13.4 | 22.575 | [6 empty] | 16
| 1 | 4 | 9.78 | 10.1504 | 13.6148 | 10.8 | 10.47 | 10.7 | 10.75 | 9.86 | 10.05 | 10.25 | 9.7558 | 11.9 | 14.8 | 13.8 | 22.535 | [6 empty] | 10.8423
| 1 | 5 | 9.61 | 9.5664 | 13.4308 | 10.7 | 10.86 | 10.15 | 10.65 | 10.05 | 10 | 10.1 | 10.041 | 12 | 13.3 | 13.8 | 20.569 | [6 empty] | 84
| 1 | 6 | 9.731 | 9.6461 | 13.504 | 10.7 | 10.5 | 10.1 | 10.75 | 10.1 | 9.89 | 10 | 9.4515 | 12.4 | 14.6 | 14.2 | 21.485
| Mean |  | 9.6355 | 9.6657 | 13.5225 | 10.8167 | 10.595 | 10.4333 | 10.6917 | 10.03 | 9.9467 | 10.15 | 9.6693 | 12.1667 | 14.8333 | 13.8 | 21.6862
| Median |  | 9.6705 | 9.6063 | 13.5238 | 10.75 | 10.5 | 10.45 | 10.675 | 10 | 9.97 | 10.15 | 9.6037 | 12.15 | 14.85 | 13.8 | 21.4835
| Std Dev. |  | 0.16812 | 0.25985 | 0.071849 | 0.19408 | 0.18053 | 0.29777 | 0.04916 | 0.18363 | 0.090701 | 0.11832 | 0.30323 | 0.20656 | 0.88694 | 0.28284 | 0.75988
| Rel.Std.Dev. |  | 1.745% | 2.688% | 0.5313% | 1.794% | 1.704% | 2.854% | 0.4598% | 1.831% | 0.9119% | 1.166% | 3.136% | 1.698% | 5.979% | 2.05% | 3.504%
| PDM3 |  | -11.13% | -10.85% | 24.72% | -0.236% | -2.28% | -3.772% | -1.389% | -7.492% | -8.26% | -6.385% | -10.82% | 12.22% | 36.81% | 27.28% | 100%
| Z-Score (Absolute) |  | 0.7 | 0.68 | 2.13 | 0.16 | 0 | 0.12 | 0.07 | 0.41 | 0.47 | 0.32 | 0.67 | 1.14 | 3.09 | 2.33 | 8.08
| Analytical results for Co in OREAS 242b (Certified Value 27.3 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 26.2 | 27.8418 | 25.52 | 29.7 | 29.3 | 28.2 | 29.9 | 26.7 | 27.2 | 25.7 | 27.7706 | 28.4 | 26 | 27 | 26.7 | 27.7 | 25 | 26.4 | 25.865 |  |  | 1
| 1 | 2 | 27.3 | 27.1429 | 27.5 | 29 | 28.4 | 27.3 | 30 | 27.7 | 27.9 | 26.5 | 27.0026 | 28.8 | 25.2 | 28 | 26.8 | 29.3 | 24.8 | 24.8 | 26.472 |  |  | 14
| 1 | 3 | 26.9 | 28.1389 | 26.943 | 29.4 | 28.3 | 27 | 30.5 | 27.4 | 28.3 | 26.1 | 26.8464 | 28.5 | 26.5 | 28 | 27.1 | 28.6 | 25.2 | 24.4 | 26.238 |  |  | 16
| 1 | 4 | 27 | 27.1114 | 27.031 | 29.4 | 28 | 28 | 30.3 | 26.7 | 28.7 | 26.5 | 26.7051 | 28.7 | 24.3 | 27 | 25.7 | 28.5 | 25.3 | 26.4 | 26.044 |  |  | 27.3313
| 1 | 5 | 26.7 | 26.7754 | 26.844 | 29.7 | 29 | 26.9 | 30.3 | 27 | 28.2 | 25.8 | 27.7926 | 28.2 | 23.5 | 29 | 27.2 | 28.5 | 25.1 | 24.6 | 25.901 |  |  | 85
| 1 | 6 | 28 | 28.04 | 27.406 | 28.7 | 28.2 | 27 | 30 | 27.9 | 27.7 | 26.5 | 26.7832 | 28.2 | 24.6 | 28 | 27.3 | 28.5 | 24.9 | 27 | 26.313
| Mean |  | 27.0167 | 27.5084 | 26.874 | 29.3167 | 28.5333 | 27.4 | 30.1667 | 27.2333 | 28 | 26.1833 | 27.1501 | 28.4667 | 25.0167 | 27.8333 | 26.8 | 28.5167 | 25.05 | 25.6 | 26.1388
| Median |  | 26.95 | 27.4923 | 26.987 | 29.4 | 28.35 | 27.15 | 30.15 | 27.2 | 28.05 | 26.3 | 26.9245 | 28.45 | 24.9 | 28 | 26.95 | 28.5 | 25.05 | 25.6 | 26.141
| Std Dev. |  | 0.6047 | 0.56916 | 0.71253 | 0.39707 | 0.50465 | 0.56214 | 0.23381 | 0.51251 | 0.52154 | 0.37103 | 0.49889 | 0.25033 | 1.1125 | 0.75277 | 0.58652 | 0.50761 | 0.18708 | 1.1243 | 0.24154
| Rel.Std.Dev. |  | 2.238% | 2.069% | 2.651% | 1.354% | 1.769% | 2.052% | 0.7751% | 1.882% | 1.863% | 1.417% | 1.838% | 0.8794% | 4.447% | 2.705% | 2.188% | 1.78% | 0.7468% | 4.392% | 0.9241%
| PDM3 |  | -1.151% | 0.6478% | -1.673% | 7.264% | 4.398% | 0.2512% | 10.37% | -0.3586% | 2.447% | -4.2% | -0.6632% | 4.154% | -8.469% | 1.837% | -1.944% | 4.337% | -8.347% | -6.335% | -4.363%
| Z-Score (Absolute) |  | 0.14 | 0.18 | 0.23 | 1.34 | 0.83 | 0.11 | 1.88 | 0 | 0.49 | 0.67 | 0.05 | 0.79 | 1.42 | 0.39 | 0.28 | 0.82 | 1.4 | 1.05 | 0.7
| Analytical results for Cr in OREAS 242b (Certified Value 21 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 40g | 0.5g |  |  | 2
| 1 | 1 | 20 | 21.0229 | 18.255 | 22.97 | 25 | 21 | 19 | 20 | 20 | 20 | 21 | 21.9581 | 24 | 20.6 | 22 | 20.9 | 22 | 20 | 17.669 |  |  | 1
| 1 | 2 | 21 | 20.8768 | 19.484 | 23.89 | 24 | 21 | 19 | 20 | 20 | 20 | 21 | 20.9445 | 23 | 19.2 | 23 | 21.1 | 23 | 20 | 17.812 |  |  | 30
| 1 | 3 | 21 | 21.7044 | 19.895 | 23.2 | 24 | 21 | 20 | 20 | 20 | 20 | 21 | 21.0456 | 24 | 19.3 | 23 | 21.7 | 22 | 20 | 17.813 |  |  | 16
| 1 | 4 | 21 | 20.7676 | 19.48 | 24.04 | 24 | 21 | 20 | 20 | 20 | 20 | 21 | 20.9715 | 23 | 19.3 | 23 | 19.8 | 22 | 20 | 18.185 |  |  | 21.0157
| 1 | 5 | 21 | 20.7009 | 19.347 | 23.58 | 23 | 21 | 20 | 20 | 20 | 21 | 21 | 21.6019 | 23 | 20 | 24 | 21.5 | 22 | 20 | 16.908 |  |  | 86
| 1 | 6 | 21 | 21.4919 | 19.578 | 23.67 | 23 | 21 | 20 | 20 | 21 | 20 | 21 | 20.9461 | 24 | 19.4 | 24 | 21.5 | 23 | 20 | 16.86
| Mean |  | 20.8333 | 21.0941 | 19.3398 | 23.5583 | 23.8333 | 21 | 19.6667 | 20 | 20.1667 | 20.1667 | 21 | 21.2446 | 23.5 | 19.6333 | 23.1667 | 21.0833 | 22.3333 | 20 | 17.5412
| Median |  | 21 | 20.9499 | 19.482 | 23.625 | 24 | 21 | 20 | 20 | 20 | 20 | 21 | 21.0085 | 23.5 | 19.35 | 23 | 21.3 | 22 | 20 | 17.7405
| Std Dev. |  | 0.40825 | 0.41088 | 0.56254 | 0.4073 | 0.75277 | 0 | 0.5164 | 0 | 0.40825 | 0.40825 | 0 | 0.43132 | 0.54772 | 0.55377 | 0.75277 | 0.69402 | 0.5164 | 0 | 0.53721
| Rel.Std.Dev. |  | 1.96% | 1.948% | 2.909% | 1.729% | 3.158% | 0% | 2.626% | 0% | 2.024% | 2.024% | 0% | 2.03% | 2.331% | 2.821% | 3.249% | 3.292% | 2.312% | 0% | 3.063%
| PDM3 |  | -0.8679% | 0.3728% | -7.975% | 12.1% | 13.41% | -0.07487% | -6.419% | -4.833% | -4.04% | -4.04% | -0.07487% | 1.089% | 11.82% | -6.578% | 10.23% | 0.3217% | 6.27% | -4.833% | -16.53%
| Z-Score (Absolute) |  | 0.11 | 0.06 | 1.12 | 1.73 | 1.91 | 0 | 0.9 | 0.67 | 0.56 | 0.56 | 0 | 0.16 | 1.69 | 0.92 | 1.46 | 0.06 | 0.9 | 0.67 | 2.33
| Analytical results for Cs in OREAS 242b (Certified Value 0.58 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 3
| 1 | 1 | 0.59 | 0.57853 | 0.52 | 0.57 | 0.57 | 0.6 | 0.56 | 0.54 | 0.58 | 0.56099 | 0.64 | 0.7 | 0.5 | [8 empty] | 1
| 1 | 2 | 0.63 | 0.56072 | 0.52 | 0.55 | 0.57 | 0.62 | 0.55 | 0.57 | 0.58 | 0.56869 | 0.66 | 0.7 | 0.5 | [8 empty] | 31
| 1 | 3 | 0.61 | 0.57036 | 0.53 | 0.53 | 0.57 | 0.61 | 0.56 | 0.54 | 0.59 | 0.56235 | 0.64 | 0.7 | 0.6 | [8 empty] | 16
| 1 | 4 | 0.62 | 0.61503 | 0.52 | 0.53 | 0.59 | 0.63 | 0.55 | 0.58 | 0.6 | 0.54973 | 0.62 | 0.7 | 0.5 | [8 empty] | 0.57704
| 1 | 5 | 0.62 | 0.56581 | 0.51 | 0.56 | 0.56 | 0.6 | 0.56 | 0.56 | 0.59 | 0.56432 | 0.63 | 0.7 | 0.5 | [8 empty] | 87
| 1 | 6 | 0.65 | 0.59668 | 0.52 | 0.54 | 0.56 | 0.62 | 0.56 | 0.55 | 0.59 | 0.55153 | 0.62 | 0.7 | 0.5
| Mean |  | 0.62 | 0.58119 | 0.52 | 0.54667 | 0.57 | 0.61333 | 0.55667 | 0.55667 | 0.58833 | 0.5596 | 0.635 | 0.7 | 0.51667
| Median |  | 0.62 | 0.57445 | 0.52 | 0.545 | 0.57 | 0.615 | 0.56 | 0.555 | 0.59 | 0.56167 | 0.635 | 0.7 | 0.5
| Std Dev. |  | 0.02 | 0.02081 | 0.0063246 | 0.01633 | 0.010954 | 0.012111 | 0.005164 | 0.01633 | 0.0075277 | 0.0074396 | 0.015166 | 1.2162e-16 | 0.040825
| Rel.Std.Dev. |  | 3.226% | 3.581% | 1.216% | 2.987% | 1.922% | 1.975% | 0.9277% | 2.934% | 1.28% | 1.329% | 2.388% | 1.737e-14% | 7.902%
| PDM3 |  | 7.445% | 0.7187% | -9.885% | -5.264% | -1.22% | 6.289% | -3.531% | -3.531% | 1.957% | -3.022% | 10.04% | 21.31% | -10.46%
| Z-Score (Absolute) |  | 1.84 | 0.41 | 1.84 | 0.86 | 0 | 1.59 | 0.49 | 0.49 | 0.67 | 0.38 | 2.39 | NA | NA
Results from laboratories 1.21 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 242b (Certified Value 170 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 0
| 1 | 1 | 162.9 | 168.2545 | 161.579 | 173.4 | 177 | 178 | 169.5 | 171.5 | 169 | 177 | 178 | 168.7035 | 175 | 171 | 180.34 | 163 | 159.9 | 149 | 165.4 | 160 | 166.033 | 1
| 1 | 2 | 168.9 | 166.9729 | 170.929 | 180.09 | 170 | 175 | 167.5 | 172 | 168 | 176 | 177 | 170.2353 | 175 | 171 | 170.05 | 166 | 160.14 | 157 | 168.6 | 159 | 167.072 | 16
| 1 | 3 | 165.8 | 170.0549 | 166.234 | 176.85 | 170 | 177 | 168 | 171 | 166.5 | 179 | 176 | 168.1089 | 176 | 173 | 174.7 | 165 | 158.1 | 156 | 173.9 | 160 | 167.204 | 16
| 1 | 4 | 169.8 | 165.8572 | 174.288 | 175.85 | 174 | 177 | 173 | 175 | 168.5 | 176 | 177 | 172.9862 | 173 | 170 | 174.62 | 162 | 154.81 | 157 | 162.8 | 161 | 167.691 | 170.2315
| 1 | 5 | 167.1 | 164.5717 | 172.559 | 176.75 | 173 | 166 | 167 | 172.5 | 167.5 | 177.5 | 174 | 172.0263 | 175 | 169 | 178.2 | 170 | 160.06 | 158 | 168 | 163 | 167.398 | 88
| 1 | 6 | 171.9 | 170.3164 | 172.897 | 176.6 | 166 | 177 | 169.5 | 173.5 | 169.5 | 173.5 | 179 | 166.711 | 176 | 171 | 176.85 | 170 | 159.38 | 152 | 165.5 | 159 | 164.302
| Mean |  | 167.7333 | 167.6713 | 169.7477 | 176.59 | 171.6667 | 175 | 169.0833 | 172.5833 | 168.1667 | 176.5 | 176.8333 | 169.7952 | 175 | 170.8333 | 175.7933 | 166 | 158.7317 | 154.8333 | 167.3667 | 160.3333 | 166.6167
| Median |  | 168 | 167.6137 | 171.744 | 176.675 | 171.5 | 177 | 168.75 | 172.25 | 168.25 | 176.5 | 177 | 169.4694 | 175 | 171 | 175.775 | 165.5 | 159.64 | 156.5 | 166.75 | 160 | 167.138
| Std Dev. |  | 3.1778 | 2.298 | 4.8766 | 2.1479 | 3.8297 | 4.5166 | 2.1775 | 1.4634 | 1.0801 | 1.8439 | 1.7224 | 2.4042 | 1.0954 | 1.3292 | 3.5536 | 3.4059 | 2.0645 | 3.5449 | 3.8161 | 1.5055 | 1.2661
| Rel.Std.Dev. |  | 1.895% | 1.371% | 2.873% | 1.216% | 2.231% | 2.581% | 1.288% | 0.848% | 0.6423% | 1.045% | 0.974% | 1.416% | 0.626% | 0.778% | 2.021% | 2.052% | 1.301% | 2.29% | 2.28% | 0.939% | 0.7599%
| PDM3 |  | -1.468% | -1.504% | -0.2842% | 3.735% | 0.8431% | 2.801% | -0.6745% | 1.382% | -1.213% | 3.682% | 3.878% | -0.2563% | 2.801% | 0.3535% | 3.267% | -2.486% | -6.755% | -9.045% | -1.683% | -5.815% | -2.123%
| Z-Score (Absolute) |  | 0.43 | 0.45 | 0 | 1.47 | 0.41 | 1.13 | 0.14 | 0.61 | 0.34 | 1.45 | 1.53 | 0.01 | 1.13 | 0.23 | 1.3 | 0.81 | 2.37 | 3.21 | 0.51 | 2.03 | 0.67
| Analytical results for Dy in OREAS 242b (Indicative Value 2.01 ppm) | [22 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 2
| 1 | 1 | 1.887 | 1.9708 | 3.5355 | 2.4 | 1.77 | 0.714 | [15 empty] | 1
| 1 | 2 | 1.95 | 1.8953 | 3.5547 | 2.3 | 1.81 | 0.78 | [15 empty] | 4
| 1 | 3 | 1.842 | 2.0163 | 3.6132 | 2.3 | 1.85 | 0.923 | [15 empty] | 16
| 1 | 4 | 1.901 | 2.0549 | 3.5859 | 2.4 | 1.9 | 0.489 | [15 empty] | 2.0148
| 1 | 5 | 1.939 | 1.951 | 3.5427 | 2.3 | 1.91 | 0.832 | [15 empty] | 10
| 1 | 6 | 1.981 | 1.988 | 3.536 | 2.3 | 1.74 | 0.433
| Mean |  | 1.9167 | 1.9794 | 3.5613 | 2.3333 | 1.83 | 0.69517
| Median |  | 1.92 | 1.9794 | 3.5487 | 2.3 | 1.83 | 0.747
| Std Dev. |  | 0.049899 | 0.054922 | 0.031622 | 0.05164 | 0.068993 | 0.19464
| Rel.Std.Dev. |  | 2.603% | 2.775% | 0.8879% | 2.213% | 3.77% | 28%
| PDM3 |  | -4.873% | -1.76% | 76.76% | 15.81% | -9.174% | -65.5%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 4.32 | 1.03 | 0.32 | 3.36
| Analytical results for Er in OREAS 242b (Indicative Value 1.15 ppm) | [22 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.959 | 1.1493 | 1.9955 | 1.5 | 1.12 | [16 empty] | 1
| 1 | 2 | 1.169 | 1.1258 | 2.0197 | 1.5 | 1.09 | [16 empty] | 5
| 1 | 3 | 1.105 | 1.1692 | 2.0676 | 1.4 | 1.12 | [16 empty] | 16
| 1 | 4 | 1.159 | 1.1817 | 1.9856 | 1.4 | 1.18 | [16 empty] | 1.15
| 1 | 5 | 1.218 | 1.1794 | 2.0717 | 1.4 | 1.19 | [16 empty] | 11
| 1 | 6 | 1.16 | 1.182 | 2.0655 | 1.4 | 1.04
| Mean |  | 1.1283 | 1.1646 | 2.0343 | 1.4333 | 1.1233
| Median |  | 1.1595 | 1.1743 | 2.0426 | 1.4 | 1.12
| Std Dev. |  | 0.090394 | 0.022678 | 0.038909 | 0.05164 | 0.056095
| Rel.Std.Dev. |  | 8.011% | 1.947% | 1.913% | 3.603% | 4.994%
| PDM3 |  | -1.887% | 1.264% | 76.89% | 24.63% | -2.322%
| Z-Score (Absolute) |  | 0.59 | 0 | 14.22 | 4.4 | 0.67
| Analytical results for Eu in OREAS 242b (Certified Value 0.44 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 2
| 1 | 1 | 0.368 | 0.37358 | 0.57652 | 0.5 | 0.298 | 0.531 | [15 empty] | 1
| 1 | 2 | 0.384 | 0.36191 | 0.61988 | 0.5 | 0.302 | 0.5 | [15 empty] | 35
| 1 | 3 | 0.369 | 0.37475 | 0.57116 | 0.5 | 0.328 | 0.534 | [15 empty] | 16
| 1 | 4 | 0.383 | 0.39238 | 0.5506 | 0.5 | 0.32 | 0.552 | [15 empty] | 0.44474
| 1 | 5 | 0.397 | 0.37573 | 0.58732 | 0.5 | 0.32 | 0.485 | [15 empty] | 89
| 1 | 6 | 0.38 | 0.37529 | 0.60064 | 0.5 | 0.3 | 0.52
| Mean |  | 0.38017 | 0.37561 | 0.58435 | 0.5 | 0.31133 | 0.52033
| Median |  | 0.3815 | 0.37502 | 0.58192 | 0.5 | 0.311 | 0.5255
| Std Dev. |  | 0.010759 | 0.0097352 | 0.024128 | 0 | 0.012817 | 0.024353
| Rel.Std.Dev. |  | 2.83% | 2.592% | 4.129% | 0% | 4.117% | 4.68%
| PDM3 |  | -14.52% | -15.54% | 31.39% | 12.43% | -30% | 17%
| Z-Score (Absolute) |  | 0.56 | 0.6 | 1.34 | 0.56 | 1.2 | 0.75
| Analytical results for Fe in OREAS 242b (Certified Value 5.29 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 5.32 | 5.2454 | 4.9572 | 5.78 | 5.23 | 5.14 | 5.22 | 5.16 | 5.3 | 5.19 | 5.797 | 5.0573 | 5.9 | 4.79 | 5.35 | 4.7 | 5.47 | 5.09 | 5.31 | 5.3306 |  | 1
| 1 | 2 | 5.57 | 5.2215 | 5.1563 | 5.71 | 5.21 | 5.07 | 5.23 | 5.22 | 5.35 | 5.14 | 5.3663 | 5.0866 | 5.79 | 4.56 | 5.5 | 4.73 | 5.59 | 5.12 | 5.38 | 5.3198 |  | #N/A
| 1 | 3 | 5.48 | 5.3803 | 5.1966 | 5.76 | 5.26 | 5.1 | 5.19 | 5.16 | 5.37 | 5.18 | 5.3621 | 5.1304 | 5.81 | 4.57 | 5.48 | 4.69 | 5.46 | 5.13 | 5.3 | 5.3824 |  | 16
| 1 | 4 | 5.58 | 5.1706 | 5.1513 | 5.76 | 5.21 | 5.24 | 5.28 | 5.14 | 5.3 | 5.15 | 5.4445 | 5.0497 | 5.71 | 4.63 | 5.44 | 4.54 | 5.49 | 5.14 | 5.33 | 5.4433 |  | 5.2873
| 1 | 5 | 5.53 | 5.1744 | 5.1761 | 5.73 | 5.21 | 5.13 | 5.25 | 5.23 | 5.38 | 5.13 | 5.5391 | 5.0874 | 5.78 | 4.77 | 5.61 | 4.75 | 5.42 | 5.12 | 5.33 | 5.2753 |  | 90
| 1 | 6 | 5.57 | 5.3646 | 5.1912 | 5.68 | 5.19 | 5.15 | 5.2 | 5.23 | 5.3 | 5.22 | 5.7068 | 5.1634 | 5.85 | 4.66 | 5.63 | 4.73 | 5.43 | 5.07 | 5.32 | 5.3002
| Mean |  | 5.5083 | 5.2595 | 5.1381 | 5.7367 | 5.2183 | 5.1383 | 5.2283 | 5.19 | 5.3333 | 5.1683 | 5.536 | 5.0958 | 5.8067 | 4.6633 | 5.5017 | 4.69 | 5.4767 | 5.1117 | 5.3283 | 5.342
| Median |  | 5.55 | 5.2335 | 5.1662 | 5.745 | 5.21 | 5.135 | 5.225 | 5.19 | 5.325 | 5.165 | 5.4918 | 5.087 | 5.8 | 4.645 | 5.49 | 4.715 | 5.465 | 5.12 | 5.325 | 5.3252
| Std Dev. |  | 0.099482 | 0.092106 | 0.090437 | 0.037238 | 0.024014 | 0.057764 | 0.033116 | 0.040988 | 0.037771 | 0.034303 | 0.18153 | 0.043657 | 0.064704 | 0.097912 | 0.10534 | 0.076681 | 0.06121 | 0.026394 | 0.027869 | 0.061178
| Rel.Std.Dev. |  | 1.806% | 1.751% | 1.76% | 0.6491% | 0.4602% | 1.124% | 0.6334% | 0.7897% | 0.7082% | 0.6637% | 3.279% | 0.8567% | 1.114% | 2.1% | 1.915% | 1.635% | 1.118% | 0.5164% | 0.523% | 1.145%
| PDM3 |  | 4.181% | -0.5255% | -2.821% | 8.5% | -1.304% | -2.817% | -1.115% | -1.84% | 0.8712% | -2.249% | 4.704% | -3.621% | 9.824% | -11.8% | 4.055% | -11.3% | 3.582% | -3.321% | 0.7767% | 1.035%
| Z-Score (Absolute) |  | 1.5 | 0.09 | 0.6 | 2.79 | 0.14 | 0.6 | 0.09 | 0.31 | 0.51 | 0.43 | 1.65 | 0.84 | 3.19 | 3.29 | 1.46 | 3.14 | 1.32 | 0.75 | 0.48 | 0.56
| Analytical results for Ga in OREAS 242b (Certified Value 9.5 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | [6 empty] | 3
| 1 | 1 | 8.65 | 9.7138 | 10 | 9.63 | 9.73 | 10 | 9.15 | 9.46 | 9.43 | 10.0341 | 11.7 | 9.5 | 8.2 | 10.3 | 14.7 | [6 empty] | 1
| 1 | 2 | 9.23 | 9.7457 | 9.84 | 9.47 | 9.56 | 9.93 | 9.19 | 9.36 | 9.5 | 9.688 | 11.6 | 8.9 | 8.3 | 10.3 | 15.8 | [6 empty] | 36
| 1 | 3 | 9.1 | 9.9229 | 9.92 | 9.66 | 9.38 | 9.97 | 9.14 | 9.51 | 9.35 | 9.8065 | 11.5 | 9.5 | 8.4 | 10.1 | 15.1 | [6 empty] | 16
| 1 | 4 | 9.19 | 9.3519 | 10 | 9.67 | 9.88 | 10.3 | 8.87 | 9.59 | 9.62 | 10.3787 | 11.4 | 8.7 | 7.9 | 10.3 | 16 | [6 empty] | 9.4956
| 1 | 5 | 8.86 | 9.2544 | 9.91 | 9.79 | 9.22 | 9.82 | 9.12 | 9.72 | 9.09 | 10.3504 | 11.5 | 8.8 | 8.5 | 10 | 14.8 | [6 empty] | 91
| 1 | 6 | 9.35 | 9.7277 | 9.76 | 9.47 | 9.37 | 10.2 | 9.34 | 9.52 | 9.63 | 9.8397 | 11.6 | 9.4 | 8.3 | 9.99 | 16.1
| Mean |  | 9.0633 | 9.6194 | 9.905 | 9.615 | 9.5233 | 10.0367 | 9.135 | 9.5267 | 9.4367 | 10.0162 | 11.55 | 9.1333 | 8.2667 | 10.165 | 15.4167
| Median |  | 9.145 | 9.7207 | 9.915 | 9.645 | 9.47 | 9.985 | 9.145 | 9.515 | 9.465 | 9.9369 | 11.55 | 9.15 | 8.3 | 10.2 | 15.45
| Std Dev. |  | 0.26059 | 0.25826 | 0.093327 | 0.12486 | 0.24792 | 0.17896 | 0.15215 | 0.1216 | 0.20136 | 0.29199 | 0.10488 | 0.37238 | 0.20656 | 0.15281 | 0.62423
| Rel.Std.Dev. |  | 2.875% | 2.685% | 0.9422% | 1.299% | 2.603% | 1.783% | 1.666% | 1.276% | 2.134% | 2.915% | 0.9081% | 4.077% | 2.499% | 1.503% | 4.049%
| PDM3 |  | -4.552% | 1.304% | 4.312% | 1.258% | 0.2925% | 5.699% | -3.797% | 0.3276% | -0.6202% | 5.483% | 21.64% | -3.815% | -12.94% | 7.05% | 62.36%
| Z-Score (Absolute) |  | 0.88 | 0.01 | 0.46 | 0 | 0.15 | 0.67 | 0.77 | 0.14 | 0.29 | 0.64 | 3.09 | 0.77 | 2.16 | 0.88 | 9.28
| Analytical results for Gd in OREAS 242b (Indicative Value 1.82 ppm) | [22 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 1.707 | 1.7339 | 2 | 1.69 | [17 empty] | 1
| 1 | 2 | 1.819 | 1.6929 | 2 | 1.77 | [17 empty] | 6
| 1 | 3 | 1.748 | 1.8084 | 2.1 | 1.65 | [17 empty] | 16
| 1 | 4 | 1.789 | 1.9098 | 2 | 1.66 | [17 empty] | 1.8156
| 1 | 5 | 1.758 | 1.7802 | 2 | 1.73 | [17 empty] | 12
| 1 | 6 | 1.86 | 1.7681 | 1.9 | 1.7
| Mean |  | 1.7802 | 1.7822 | 2 | 1.7
| Median |  | 1.7735 | 1.7742 | 2 | 1.695
| Std Dev. |  | 0.054477 | 0.074165 | 0.063246 | 0.044721
| Rel.Std.Dev. |  | 3.06% | 4.161% | 3.162% | 2.631%
| PDM3 |  | -1.951% | -1.838% | 10.16% | -6.367%
| Z-Score (Absolute) |  | 0.02 | 0.02 | 3.59 | 1.33
| Analytical results for Ge in OREAS 242b (Certified Value 0.14 ppm) | [22 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.21 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.1g | [12 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.15 | 0.11 | 0.15 | 0.17 | 0.13 | 0.2 | 0.15 | [12 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.15 | 0.11 | 0.16 | 0.18 | 0.13 | 0.2 | 0.14 | [12 empty] | 38
| 1 | 3 | < 0.2 | < 0.1 | 0.14 | 0.11 | 0.14 | 0.17 | 0.12 | 0.22 | 0.14 | [12 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.13 | 0.11 | 0.15 | 0.18 | 0.12 | 0.21 | 0.13 | [12 empty] | 0.13917
| 1 | 5 | < 0.2 | < 0.1 | 0.15 | 0.11 | 0.14 | 0.16 | 0.11 | 0.21 | 0.14 | [12 empty] | 92
| 1 | 6 | < 0.2 | < 0.1 | 0.13 | 0.11 | 0.16 | 0.17 | 0.12 | 0.21 | 0.14
| Mean |  |  |  | 0.14167 | 0.11 | 0.15 | 0.17167 | 0.12167 | 0.20833 | 0.14
| Median |  |  |  | 0.145 | 0.11 | 0.15 | 0.17 | 0.12 | 0.21 | 0.14
| Std Dev. |  |  |  | 0.0098319 | 0 | 0.0089443 | 0.0075277 | 0.0075277 | 0.0075277 | 0.0063246
| Rel.Std.Dev. |  |  |  | 6.94% | 0% | 5.963% | 4.385% | 6.187% | 3.613% | 4.518%
| PDM3 |  |  |  | 1.796% | -20.96% | 7.784% | 23.35% | -12.57% | 49.7% | 0.5988%
| Z-Score (Absolute) |  | 0.9 | 2.02 | 0.04 | 0.67 | 0.22 | 0.71 | 0.41 | 1.54 | 0
| Analytical results for Hf in OREAS 242b (Certified Value 0.47 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 3
| 1 | 1 | 0.43 | 0.43571 | 0.3 | 0.52 | 0.45 | 0.53 | 0.41 | 0.42 | 0.45 | 0.54033 | 0.59 | 0.62 | [9 empty] | 1
| 1 | 2 | 0.51 | 0.41143 | 0.3 | 0.49 | 0.45 | 0.53 | 0.43 | 0.45 | 0.45 | 0.52851 | 0.6 | 0.62 | [9 empty] | 22
| 1 | 3 | 0.44 | 0.42377 | 0.4 | 0.5 | 0.48 | 0.52 | 0.44 | 0.46 | 0.46 | 0.51794 | 0.59 | 0.62 | [9 empty] | 16
| 1 | 4 | 0.48 | 0.44513 | 0.3 | 0.44 | 0.43 | 0.53 | 0.45 | 0.45 | 0.46 | 0.5352 | 0.63 | 0.65 | [9 empty] | 0.47032
| 1 | 5 | 0.49 | 0.43426 | 0.3 | 0.52 | 0.44 | 0.56 | 0.44 | 0.43 | 0.47 | 0.52222 | 0.61 | 0.6 | [9 empty] | 93
| 1 | 6 | 0.47 | 0.44661 | 0.3 | 0.46 | 0.43 | 0.49 | 0.44 | 0.44 | 0.44 | 0.54606 | 0.62 | 0.62
| Mean |  | 0.47 | 0.43282 | 0.31667 | 0.48833 | 0.44667 | 0.52667 | 0.435 | 0.44167 | 0.455 | 0.53171 | 0.60667 | 0.62167
| Median |  | 0.475 | 0.43499 | 0.3 | 0.495 | 0.445 | 0.53 | 0.44 | 0.445 | 0.455 | 0.53186 | 0.605 | 0.62
| Std Dev. |  | 0.030332 | 0.013354 | 0.040825 | 0.032506 | 0.018619 | 0.022509 | 0.013784 | 0.01472 | 0.010488 | 0.010793 | 0.01633 | 0.016021
| Rel.Std.Dev. |  | 6.454% | 3.085% | 12.89% | 6.657% | 4.168% | 4.274% | 3.169% | 3.333% | 2.305% | 2.03% | 2.692% | 2.577%
| PDM3 |  | -0.06763% | -7.973% | -32.67% | 3.83% | -5.029% | 11.98% | -7.509% | -6.092% | -3.257% | 13.05% | 28.99% | 32.18%
| Z-Score (Absolute) |  | 0 | 0.72 | NA | 0.35 | 0.45 | 1.09 | 0.67 | 0.55 | 0.29 | 1.19 | 2.63 | 2.92
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 242b (Certified Value 0.07 ppm) | [22 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.1g | 01g | 40g | [6 empty] | 3
| 1 | 1 | < 1 | 0.09 | 0.34 | 0.06 | 0.07 | 0.06 | 0.06 | 0.09 | < 0.05 | 0.14 | 0.066 | 0.066 | 0.11 | < 5 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 1 | 0.11 | 0.32 | 0.06 | 0.05 | 0.06 | 0.08 | 0.08 | < 0.05 | 0.14 | 0.059 | 0.056 | 0.11 | < 5 | < 0.5 | [6 empty] | 2
| 1 | 3 | < 1 | 0.09 | 0.31 | 0.06 | 0.06 | 0.06 | 0.07 | 0.09 | < 0.05 | 0.15 | 0.08 | 0.062 | 0.12 | < 5 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 1 | 0.09 | 0.33 | 0.07 | 0.09 | 0.06 | 0.07 | 0.09 | < 0.05 | 0.14 | 0.06 | 0.059 | 0.13 | < 5 | < 0.5 | [6 empty] | 0.069917
| 1 | 5 | < 1 | 0.08 | 0.3 | 0.06 | 0.05 | 0.06 | 0.06 | 0.09 | < 0.05 | 0.15 | 0.085 | 0.06 | 0.12 | < 5 | < 0.5 | [6 empty] | 94
| 1 | 6 | < 1 | 0.08 | 0.34 | 0.06 | 0.05 | 0.07 | 0.05 | 0.1 | < 0.05 | 0.14 | 0.06 | 0.063 | 0.11 | < 5 | < 0.5
| Mean |  |  | 0.09 | 0.32333 | 0.061667 | 0.061667 | 0.061667 | 0.065 | 0.09 |  | 0.14333 | 0.068333 | 0.061 | 0.11667
| Median |  |  | 0.09 | 0.325 | 0.06 | 0.055 | 0.06 | 0.065 | 0.09 |  | 0.14 | 0.063 | 0.061 | 0.115
| Std Dev. |  |  | 0.010954 | 0.01633 | 0.0040825 | 0.016021 | 0.0040825 | 0.010488 | 0.0063246 |  | 0.005164 | 0.011361 | 0.0034641 | 0.008165
| Rel.Std.Dev. |  |  | 12.17% | 5.05% | 6.62% | 25.98% | 6.62% | 16.14% | 7.027% |  | 3.603% | 16.63% | 5.679% | 6.999%
| PDM3 |  |  | 28.72% | 362.5% | -11.8% | -11.8% | -11.8% | -7.032% | 28.72% |  | 105% | -2.265% | -12.75% | 66.87%
| Z-Score (Absolute) |  | 9.76 | 0 | 5.55 | 0.67 | 0.67 | 0.67 | 0.59 | 0 | 1.55 | 1.27 | 0.52 | 0.69 | 0.63 | 57.36 | 3.81
| Analytical results for Ho in OREAS 242b (Certified Value 0.44 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.39 | 0.39972 | 0.5554 | 0.5 | 0.37 | [16 empty] | 1
| 1 | 2 | 0.396 | 0.38175 | 0.5596 | 0.5 | 0.395 | [16 empty] | 23
| 1 | 3 | 0.379 | 0.39544 | 0.5024 | 0.5 | 0.385 | [16 empty] | 16
| 1 | 4 | 0.407 | 0.39634 | 0.51796 | 0.5 | 0.365 | [16 empty] | 0.43994
| 1 | 5 | 0.405 | 0.37942 | 0.553 | 0.4 | 0.41 | [16 empty] | 95
| 1 | 6 | 0.417 | 0.39544 | 0.56764 | 0.5 | 0.375
| Mean |  | 0.399 | 0.39135 | 0.54267 | 0.48333 | 0.38333
| Median |  | 0.4005 | 0.39544 | 0.5542 | 0.5 | 0.38
| Std Dev. |  | 0.01352 | 0.0085183 | 0.02612 | 0.040825 | 0.016931
| Rel.Std.Dev. |  | 3.389% | 2.177% | 4.813% | 8.447% | 4.417%
| PDM3 |  | -9.305% | -11.04% | 23.35% | 9.864% | -12.87%
| Z-Score (Absolute) |  | 0 | 0.33 | 6.18 | 3.63 | 0.67
| Analytical results for In in OREAS 242b (Certified Value 0.032 ppm) | [22 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [9 empty] | 3
| 1 | 1 | 0.03 | 0.05 | 0.095 | 0.034 | 0.032 | 0.031 | 0.032 | 0.032 | < 0.05 | < 0.05 | 0.06 | < 0.5 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.05 | 0.109 | 0.031 | 0.034 | 0.032 | 0.031 | 0.031 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 24
| 1 | 3 | 0.03 | 0.04 | 0.112 | 0.031 | 0.037 | 0.032 | 0.034 | 0.032 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 16
| 1 | 4 | 0.04 | 0.04 | 0.092 | 0.033 | 0.034 | 0.03 | 0.032 | 0.031 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 0.032
| 1 | 5 | 0.03 | 0.04 | 0.105 | 0.033 | 0.034 | 0.031 | 0.033 | 0.031 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 96
| 1 | 6 | 0.03 | 0.03 | 0.095 | 0.034 | 0.033 | 0.033 | 0.033 | 0.031 | < 0.05 | < 0.05 | 0.05 | < 0.5
| Mean |  | 0.031667 | 0.041667 | 0.10133 | 0.032667 | 0.034 | 0.0315 | 0.0325 | 0.031333 |  |  | 0.045
| Median |  | 0.03 | 0.04 | 0.1 | 0.033 | 0.034 | 0.0315 | 0.0325 | 0.031 |  |  | 0.04
| Std Dev. |  | 0.0040825 | 0.0075277 | 0.0084063 | 0.0013663 | 0.0016733 | 0.0010488 | 0.0010488 | 0.0005164 |  |  | 0.0083666
| Rel.Std.Dev. |  | 12.89% | 18.07% | 8.296% | 4.182% | 4.922% | 3.33% | 3.227% | 1.648% |  |  | 18.59%
| PDM3 |  | -1.042% | 30.21% | 216.7% | 2.083% | 6.25% | -1.562% | 1.562% | -2.083% |  |  | 40.63%
| Z-Score (Absolute) |  | 0.14 | 1.36 | 10.3 | 0.01 | 0.21 | 0.16 | 0.01 | 0.19 | 1.14 | 1.14 | 1.86 | 32.58
| Analytical results for K in OREAS 242b (Certified Value 0.108 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.1033 | 0.10915 | 0.098459 | 0.1 | 0.11 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.10346 | 0.138 | 0.11 | 0.1205 | 0.11 | 0.14 | 0.12 | 0.09 | 0.11402 |  |  | 1
| 1 | 2 | 0.1086 | 0.11208 | 0.099884 | 0.1 | 0.11 | 0.11 | 0.11 | 0.1 | 0.11 | 0.1 | 0.10012 | 0.137 | 0.1 | 0.1239 | 0.11 | 0.133 | 0.12 | 0.1 | 0.11471 |  |  | #N/A
| 1 | 3 | 0.1063 | 0.11394 | 0.10473 | 0.1 | 0.11 | 0.1 | 0.11 | 0.1 | 0.11 | 0.11 | 0.10078 | 0.141 | 0.11 | 0.1227 | 0.11 | 0.124 | 0.12 | 0.09 | 0.11487 |  |  | 16
| 1 | 4 | 0.1088 | 0.11253 | 0.09855 | 0.1 | 0.11 | 0.11 | 0.11 | 0.1 | 0.11 | 0.1 | 0.10391 | 0.143 | 0.1 | 0.1205 | 0.1 | 0.121 | 0.12 | 0.1 | 0.11744 |  |  | 0.1085
| 1 | 5 | 0.1078 | 0.11417 | 0.10552 | 0.1 | 0.11 | 0.1 | 0.11 | 0.1 | 0.11 | 0.11 | 0.10476 | 0.136 | 0.1 | 0.1245 | 0.11 | 0.114 | 0.12 | 0.1 | 0.11368 |  |  | 97
| 1 | 6 | 0.1075 | 0.11161 | 0.10153 | 0.1 | 0.11 | 0.11 | 0.11 | 0.1 | 0.11 | 0.11 | 0.10059 | 0.138 | 0.1 | 0.1248 | 0.11 | 0.132 | 0.12 | 0.1 | 0.11056
| Mean |  | 0.10705 | 0.11225 | 0.10145 | 0.1 | 0.11 | 0.105 | 0.11 | 0.1 | 0.10833 | 0.105 | 0.10227 | 0.13883 | 0.10333 | 0.12282 | 0.10833 | 0.12733 | 0.12 | 0.096667 | 0.11422
| Median |  | 0.10765 | 0.1123 | 0.10071 | 0.1 | 0.11 | 0.105 | 0.11 | 0.1 | 0.11 | 0.105 | 0.10212 | 0.138 | 0.1 | 0.1233 | 0.11 | 0.128 | 0.12 | 0.1 | 0.11437
| Std Dev. |  | 0.0020423 | 0.0018271 | 0.0030713 | 1.5202e-17 | 0 | 0.0054772 | 0 | 1.5202e-17 | 0.0040825 | 0.0054772 | 0.001999 | 0.0026394 | 0.005164 | 0.0019333 | 0.0040825 | 0.0094163 | 0 | 0.005164 | 0.0022238
| Rel.Std.Dev. |  | 1.908% | 1.628% | 3.028% | 1.52e-14% | 0% | 5.216% | 0% | 1.52e-14% | 3.768% | 5.216% | 1.955% | 1.901% | 4.997% | 1.574% | 3.768% | 7.395% | 0% | 5.342% | 1.947%
| PDM3 |  | -1.334% | 3.455% | -6.5% | -7.832% | 1.385% | -3.223% | 1.385% | -7.832% | -0.151% | -3.223% | -5.742% | 27.96% | -4.759% | 13.2% | -0.151% | 17.36% | 10.6% | -10.9% | 5.271%
| Z-Score (Absolute) |  | 0.17 | 0.53 | 0.93 | NA | 0.22 | 0.45 | 0.22 | NA | 0 | 0.45 | 0.82 | 4.11 | 0.67 | 1.95 | 0 | 2.56 | 1.57 | 1.57 | 0.79
Results from laboratories 1.07 and 1.11 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for La in OREAS 242b (Certified Value 3.92 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 2
| 1 | 1 | 3.64 | 3.7366 | 6.4189 | 4.6 | 4.2 | 4.1 | 4.3 | 3.7 | 3.9 | 3.9 | 3.935 | 4.95 | 3.9 | 3.5 | 5.64 | 3.1 | 5.2 | 3.607 |  |  |  | 1
| 1 | 2 | 3.776 | 3.666 | 6.4542 | 4.4 | 4.1 | 4.1 | 4.3 | 3.8 | 4 | 4 | 3.6917 | 5.06 | 3.7 | 3.5 | 5.41 | 3.1 | 5.2 | 3.434 |  |  |  | 25
| 1 | 3 | 3.628 | 3.8384 | 6.4935 | 4.4 | 4.1 | 3.9 | 4.3 | 3.7 | 4.1 | 4 | 3.6446 | 5.01 | 4.2 | 3.4 | 5.71 | 3.1 | 5.4 | 3.51 |  |  |  | 16
| 1 | 4 | 3.862 | 3.9684 | 6.4472 | 4.4 | 4.1 | 4.2 | 4.4 | 3.7 | 4 | 4.1 | 3.8162 | 5.09 | 3.4 | 3.3 | 5.26 | 3.1 | 5 | 3.64 |  |  |  | 3.9156
| 1 | 5 | 3.714 | 3.8166 | 6.459 | 4.3 | 4.3 | 3.8 | 4.2 | 3.7 | 4.1 | 4.1 | 3.9118 | 4.92 | 3.5 | 3.5 | 4.83 | 3.1 | 4.8 | 3.499 |  |  |  | 98
| 1 | 6 | 3.776 | 3.7651 | 6.4589 | 4.4 | 4.1 | 4 | 4.3 | 3.8 | 4 | 4 | 3.6597 | 5.07 | 3.6 | 3.4 | 5.36 | 3.1 | 5.2 | 3.364
| Mean |  | 3.7327 | 3.7985 | 6.4553 | 4.4167 | 4.15 | 4.0167 | 4.3 | 3.7333 | 4.0167 | 4.0167 | 3.7765 | 5.0167 | 3.7167 | 3.4333 | 5.3683 | 3.1 | 5.1333 | 3.509
| Median |  | 3.745 | 3.7909 | 6.4566 | 4.4 | 4.1 | 4.05 | 4.3 | 3.7 | 4 | 4 | 3.754 | 5.035 | 3.65 | 3.45 | 5.385 | 3.1 | 5.2 | 3.5045
| Std Dev. |  | 0.089859 | 0.10324 | 0.023986 | 0.098319 | 0.083666 | 0.1472 | 0.063246 | 0.05164 | 0.075277 | 0.075277 | 0.12901 | 0.069186 | 0.29269 | 0.08165 | 0.31429 | 0 | 0.20656 | 0.10346
| Rel.Std.Dev. |  | 2.407% | 2.718% | 0.3716% | 2.226% | 2.016% | 3.665% | 1.471% | 1.383% | 1.874% | 1.874% | 3.416% | 1.379% | 7.875% | 2.378% | 5.854% | 0% | 4.024% | 2.948%
| PDM3 |  | -4.671% | -2.989% | 64.86% | 12.8% | 5.987% | 2.582% | 9.818% | -4.654% | 2.582% | 2.582% | -3.551% | 28.12% | -5.08% | -12.32% | 37.1% | -20.83% | 31.1% | -10.38%
| Z-Score (Absolute) |  | 0.66 | 0.5 | 5.63 | 0.92 | 0.31 | 0 | 0.65 | 0.65 | 0 | 0 | 0.55 | 2.31 | 0.69 | 1.35 | 3.12 | 2.12 | 2.58 | 1.17
| Analytical results for Li in OREAS 242b (Certified Value 9.17 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 3
| 1 | 1 | 8.1 | < 10 | 10.2874 | 8.8 | 7.5 | 9.2 | 9.7 | 9.4 | 9.5 | 9.5 | 7.9237 | 9.5 | 10.9 | < 10 | 8.121 | [6 empty] | 1
| 1 | 2 | 8.4 | < 10 | 10.4649 | 8.5 | 7.9 | 9.1 | 9.7 | 9.2 | 9.4 | 9.4 | 8.1233 | 9.3 | 9.8 | < 10 | 8.852 | [6 empty] | #N/A
| 1 | 3 | 8.3 | < 10 | 10.4775 | 8.6 | 8.3 | 9.4 | 9.7 | 9.3 | 9.5 | 9.5 | 7.877 | 9.4 | 9 | < 10 | 8.906 | [6 empty] | 16
| 1 | 4 | 8.4 | < 10 | 10.5973 | 8.6 | 8 | 9.6 | 9.8 | 9.4 | 9.4 | 9.5 | 8.2104 | 9.1 | 10.7 | < 10 | 8.962 | [6 empty] | 9.1655
| 1 | 5 | 8.3 | < 10 | 10.2587 | 8.6 | 8.1 | 9.4 | 9.8 | 9.4 | 9.4 | 9.4 | 8.2894 | 9.4 | 10.9 | < 10 | 8.91 | [6 empty] | 99
| 1 | 6 | 8.6 | < 10 | 10.6844 | 8.6 | 7.8 | 9.5 | 9.6 | 9.4 | 9.3 | 9.4 | 7.9924 | 9.5 | 9.8 | < 10 | 8.72
| Mean |  | 8.35 |  | 10.4617 | 8.6167 | 7.9333 | 9.3667 | 9.7167 | 9.35 | 9.4167 | 9.45 | 8.0694 | 9.3667 | 10.1833 |  | 8.7452
| Median |  | 8.35 |  | 10.4712 | 8.6 | 7.95 | 9.4 | 9.7 | 9.4 | 9.4 | 9.45 | 8.0579 | 9.4 | 10.25 |  | 8.879
| Std Dev. |  | 0.16432 |  | 0.1672 | 0.098319 | 0.27325 | 0.18619 | 0.075277 | 0.083666 | 0.075277 | 0.054772 | 0.16451 | 0.15055 | 0.77309 |  | 0.31676
| Rel.Std.Dev. |  | 1.968% |  | 1.598% | 1.141% | 3.444% | 1.988% | 0.7747% | 0.8948% | 0.7994% | 0.5796% | 2.039% | 1.607% | 7.592% |  | 3.622%
| PDM3 |  | -8.897% |  | 14.14% | -5.988% | -13.44% | 2.195% | 6.014% | 2.013% | 2.741% | 3.104% | -11.96% | 2.195% | 11.11% |  | -4.586%
| Z-Score (Absolute) |  | 0.92 | 4 | 1.02 | 0.67 | 1.3 | 0.02 | 0.34 | 0 | 0.06 | 0.09 | 1.18 | 0.02 | 0.77 | 4 | 0.56
| Analytical results for Lu in OREAS 242b (Indicative Value 0.15 ppm) | [22 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 0.139 | 0.13675 | 0.2 | 0.125 | [17 empty] | 1
| 1 | 2 | 0.153 | 0.12834 | 0.2 | 0.13 | [17 empty] | 7
| 1 | 3 | 0.143 | 0.14012 | 0.2 | 0.135 | [17 empty] | 16
| 1 | 4 | 0.142 | 0.13959 | 0.2 | 0.125 | [17 empty] | 0.15259
| 1 | 5 | 0.146 | 0.13333 | 0.2 | 0.14 | [17 empty] | 13
| 1 | 6 | 0.149 | 0.1321 | 0.2 | 0.125
| Mean |  | 0.14533 | 0.13504 | 0.2 | 0.13
| Median |  | 0.1445 | 0.13504 | 0.2 | 0.1275
| Std Dev. |  | 0.0050859 | 0.0046008 | 3.0405e-17 | 0.0063246
| Rel.Std.Dev. |  | 3.499% | 3.407% | 1.52e-14% | 4.865%
| PDM3 |  | -4.758% | -11.5% | 31.07% | -14.81%
| Z-Score (Absolute) |  | 0.45 | 0.45 | 5.26 | 0.9
| Analytical results for Mg in OREAS 242b (Certified Value 1.56 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 1.49 | 1.5488 | 1.5385 | 1.64 | 1.51 | 1.51 | 1.55 | 1.49 | 1.57 | 1.5 | 1.5567 | 1.64 | 1.56 | 1.51 | 1.52 | 1.34 | 1.62 | 1.65 | 1.6051 |  |  | 1
| 1 | 2 | 1.56 | 1.5513 | 1.5685 | 1.6 | 1.5 | 1.5 | 1.55 | 1.52 | 1.59 | 1.49 | 1.5541 | 1.62 | 1.47 | 1.55 | 1.52 | 1.3 | 1.64 | 1.65 | 1.6159 |  |  | #N/A
| 1 | 3 | 1.54 | 1.5592 | 1.6176 | 1.61 | 1.51 | 1.51 | 1.53 | 1.5 | 1.6 | 1.5 | 1.5474 | 1.66 | 1.49 | 1.55 | 1.53 | 1.29 | 1.63 | 1.65 | 1.6065 |  |  | 16
| 1 | 4 | 1.56 | 1.5692 | 1.5601 | 1.62 | 1.51 | 1.55 | 1.56 | 1.49 | 1.58 | 1.49 | 1.5557 | 1.6 | 1.45 | 1.53 | 1.48 | 1.2 | 1.64 | 1.65 | 1.6219 |  |  | 1.5604
| 1 | 5 | 1.55 | 1.5676 | 1.6007 | 1.61 | 1.48 | 1.52 | 1.55 | 1.52 | 1.6 | 1.48 | 1.5476 | 1.61 | 1.49 | 1.58 | 1.53 | 1.24 | 1.63 | 1.65 | 1.5933 |  |  | 100
| 1 | 6 | 1.56 | 1.5694 | 1.6091 | 1.6 | 1.5 | 1.53 | 1.55 | 1.52 | 1.57 | 1.51 | 1.5518 | 1.63 | 1.44 | 1.58 | 1.54 | 1.3 | 1.62 | 1.66 | 1.5966
| Mean |  | 1.5433 | 1.5609 | 1.5824 | 1.6133 | 1.5017 | 1.52 | 1.5483 | 1.5067 | 1.585 | 1.495 | 1.5522 | 1.6267 | 1.4833 | 1.55 | 1.52 | 1.2783 | 1.63 | 1.6517 | 1.6066
| Median |  | 1.555 | 1.5634 | 1.5846 | 1.61 | 1.505 | 1.515 | 1.55 | 1.51 | 1.585 | 1.495 | 1.5529 | 1.625 | 1.48 | 1.55 | 1.525 | 1.295 | 1.63 | 1.65 | 1.6058
| Std Dev. |  | 0.027325 | 0.0092468 | 0.031306 | 0.015055 | 0.01169 | 0.017889 | 0.0098319 | 0.015055 | 0.013784 | 0.010488 | 0.0040198 | 0.021602 | 0.04274 | 0.027568 | 0.020976 | 0.049967 | 0.0089443 | 0.0040825 | 0.010973
| Rel.Std.Dev. |  | 1.771% | 0.5924% | 1.978% | 0.9332% | 0.7785% | 1.177% | 0.635% | 0.9993% | 0.8697% | 0.7015% | 0.259% | 1.328% | 2.881% | 1.779% | 1.38% | 3.909% | 0.5487% | 0.2472% | 0.683%
| PDM3 |  | -1.096% | 0.03006% | 1.409% | 3.39% | -3.766% | -2.591% | -0.7752% | -3.445% | 1.575% | -4.193% | -0.5268% | 4.245% | -4.941% | -0.6684% | -2.591% | -18.08% | 4.458% | 5.847% | 2.956%
| Z-Score (Absolute) |  | 0.1 | 0.17 | 0.5 | 0.99 | 0.75 | 0.47 | 0.03 | 0.67 | 0.54 | 0.86 | 0.03 | 1.19 | 1.04 | 0 | 0.47 | 4.23 | 1.24 | 1.58 | 0.88
| Analytical results for Mn in OREAS 242b (Certified Value 0.067 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0661 | 0.068225 | 0.062764 | 0.0757 | 0.065 | 0.0631 | 0.0653 | 0.0641 | 0.0648 | 0.0659 | 0.070203 | 0.0737 | 0.0695 | 0.0709 | 0.0677 | 0.068 | 0.06708 | 0.067 | 0.066438 |  |  | 1
| 1 | 2 | 0.0691 | 0.067645 | 0.064996 | 0.0731 | 0.0651 | 0.0627 | 0.0648 | 0.0649 | 0.0659 | 0.0653 | 0.067613 | 0.0725 | 0.0653 | 0.0729 | 0.0681 | 0.0694 | 0.06723 | 0.067 | 0.066555 |  |  | 27
| 1 | 3 | 0.0682 | 0.069912 | 0.065469 | 0.0749 | 0.0655 | 0.0635 | 0.0657 | 0.0643 | 0.0666 | 0.0658 | 0.067522 | 0.0723 | 0.0666 | 0.0724 | 0.0681 | 0.0679 | 0.06774 | 0.068 | 0.066502 |  |  | 16
| 1 | 4 | 0.0692 | 0.066431 | 0.064642 | 0.0754 | 0.0648 | 0.0654 | 0.0664 | 0.0644 | 0.0661 | 0.0652 | 0.068508 | 0.0731 | 0.0657 | 0.0716 | 0.0657 | 0.068 | 0.06795 | 0.068 | 0.067108 |  |  | 0.066646
| 1 | 5 | 0.0688 | 0.06712 | 0.066088 | 0.0749 | 0.0652 | 0.0642 | 0.0649 | 0.0659 | 0.0675 | 0.065 | 0.070339 | 0.0724 | 0.0681 | 0.0738 | 0.0685 | 0.069 | 0.06766 | 0.068 | 0.065552 |  |  | 101
| 1 | 6 | 0.0684 | 0.069884 | 0.064976 | 0.0742 | 0.0653 | 0.0636 | 0.0647 | 0.0654 | 0.0656 | 0.0667 | 0.067099 | 0.0733 | 0.0668 | 0.0743 | 0.0686 | 0.0679 | 0.06709 | 0.067 | 0.065769
| Mean |  | 0.0683 | 0.068203 | 0.064822 | 0.0747 | 0.06515 | 0.06375 | 0.0653 | 0.064833 | 0.066083 | 0.06565 | 0.068547 | 0.072883 | 0.067 | 0.07265 | 0.067783 | 0.068367 | 0.067458 | 0.0675 | 0.066321
| Median |  | 0.0686 | 0.067935 | 0.064986 | 0.0749 | 0.06515 | 0.06355 | 0.0651 | 0.06465 | 0.066 | 0.06555 | 0.06806 | 0.0728 | 0.0667 | 0.07265 | 0.0681 | 0.068 | 0.067445 | 0.0675 | 0.06647
| Std Dev. |  | 0.0011454 | 0.0014403 | 0.0011264 | 0.00093595 | 0.0002429 | 0.00095237 | 0.00065422 | 0.00070333 | 0.00091524 | 0.00062209 | 0.0014124 | 0.00056716 | 0.0015646 | 0.0012911 | 0.0010704 | 0.00065929 | 0.00037221 | 0.00054772 | 0.00056865
| Rel.Std.Dev. |  | 1.677% | 2.112% | 1.738% | 1.253% | 0.3728% | 1.494% | 1.002% | 1.085% | 1.385% | 0.9476% | 2.061% | 0.7782% | 2.335% | 1.777% | 1.579% | 0.9643% | 0.5518% | 0.8114% | 0.8574%
| PDM3 |  | 2.482% | 2.336% | -2.736% | 12.08% | -2.245% | -4.345% | -2.02% | -2.72% | -0.8443% | -1.495% | 2.853% | 9.359% | 0.5311% | 9.009% | 1.706% | 2.582% | 1.219% | 1.281% | -0.488%
| Z-Score (Absolute) |  | 0.41 | 0.37 | 1.29 | 3.55 | 1.13 | 1.82 | 1.06 | 1.29 | 0.67 | 0.89 | 0.53 | 2.66 | 0.22 | 2.55 | 0.16 | 0.45 | 0 | 0.02 | 0.56
| Analytical results for Mo in OREAS 242b (Certified Value 7.85 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 7.8 | 7.5618 | 8.6165 | 7.91 | 8.18 | 8.01 | 7.97 | 7.65 | 7.69 | 8.79 | 6.8299 | 8.2 | 7.9 | 6.88 | 8 | 7.59 | 7.5 | 12.4 | 8.6 | 7.736 |  | 1
| 1 | 2 | 8.2 | 7.3725 | 8.256 | 7.69 | 7.87 | 7.88 | 8.04 | 7.87 | 7.88 | 8.39 | 6.8861 | 8.3 | 8.1 | 6.88 | 8 | 7.58 | 7.5 | 12.2 | 8.4 | 7.598 |  | 28
| 1 | 3 | 7.9 | 7.5156 | 8.5527 | 7.79 | 8.12 | 7.73 | 8.12 | 7.65 | 8.12 | 8.66 | 6.8072 | 8.6 | 7.9 | 7.36 | 8 | 7.65 | 7.3 | 11.8 | 8.4 | 7.974 |  | 16
| 1 | 4 | 8.2 | 7.6149 | 8.7481 | 7.74 | 7.99 | 8.01 | 8.04 | 7.54 | 7.9 | 8.47 | 6.658 | 8.2 | 7.9 | 6.74 | 7 | 7.36 | 7.3 | 11.8 | 8.8 | 7.896 |  | 7.8494
| 1 | 5 | 8 | 7.4914 | 8.7522 | 7.77 | 8.19 | 7.6 | 7.93 | 7.72 | 8.13 | 8.66 | 6.928 | 8.4 | 7.9 | 6.56 | 8 | 7.72 | 7.4 | 12.4 | 8.2 | 7.723 |  | 102
| 1 | 6 | 8.2 | 7.5779 | 8.5557 | 7.68 | 8.05 | 7.69 | 8.02 | 7.85 | 7.97 | 8.84 | 6.6674 | 8.2 | 7.8 | 6.89 | 8 | 7.56 | 7.5 | 11.3 | 8.8 | 8.02
| Mean |  | 8.05 | 7.5224 | 8.5802 | 7.7633 | 8.0667 | 7.82 | 8.02 | 7.7133 | 7.9483 | 8.635 | 6.7961 | 8.3167 | 7.9167 | 6.885 | 7.8333 | 7.5767 | 7.4167 | 11.9833 | 8.5333 | 7.8245
| Median |  | 8.1 | 7.5387 | 8.5861 | 7.755 | 8.085 | 7.805 | 8.03 | 7.685 | 7.935 | 8.66 | 6.8185 | 8.25 | 7.9 | 6.88 | 8 | 7.585 | 7.45 | 12 | 8.5 | 7.816
| Std Dev. |  | 0.17607 | 0.085624 | 0.18196 | 0.083825 | 0.12307 | 0.17274 | 0.065422 | 0.12754 | 0.16534 | 0.17581 | 0.11171 | 0.16021 | 0.098319 | 0.26546 | 0.40825 | 0.12111 | 0.098319 | 0.43089 | 0.24221 | 0.16439
| Rel.Std.Dev. |  | 2.187% | 1.138% | 2.121% | 1.08% | 1.526% | 2.209% | 0.8157% | 1.654% | 2.08% | 2.036% | 1.644% | 1.926% | 1.242% | 3.856% | 5.212% | 1.598% | 1.326% | 3.596% | 2.838% | 2.101%
| PDM3 |  | 2.555% | -4.167% | 9.31% | -1.097% | 2.767% | -0.375% | 2.173% | -1.734% | 1.26% | 10.01% | -13.42% | 5.952% | 0.8565% | -12.29% | -0.2051% | -3.475% | -5.513% | 52.66% | 8.713% | -0.3177%
| Z-Score (Absolute) |  | 0.26 | 0.78 | 1.32 | 0.3 | 0.3 | 0.19 | 0.2 | 0.4 | 0.06 | 1.42 | 2.22 | 0.79 | 0 | 2.05 | NA | 0.67 | 0.99 | 8.07 | 1.22 | 0.18
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 242b (Certified Value 0.117 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.09 | 0.12069 | 0.12458 | 0.109 | 0.11 | 0.12 | 0.13 | 0.12 | 0.12 | 0.11 | 0.11839 | 0.14 | 0.123 | 0.13 | 0.121 | 0.101 | 0.09 | 0.12 | 0.10405 |  |  | 1
| 1 | 2 | 0.12 | 0.12192 | 0.12976 | 0.11 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.11 | 0.11487 | 0.14 | 0.115 | 0.13 | 0.121 | 0.0977 | 0.09 | 0.11 | 0.11163 |  |  | #N/A
| 1 | 3 | 0.12 | 0.12451 | 0.13795 | 0.105 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.11 | 0.11694 | 0.15 | 0.119 | 0.14 | 0.12 | 0.096 | 0.09 | 0.11 | 0.10696 |  |  | 16
| 1 | 4 | 0.12 | 0.12185 | 0.13244 | 0.108 | 0.12 | 0.12 | 0.13 | 0.11 | 0.12 | 0.1 | 0.11538 | 0.14 | 0.114 | 0.13 | 0.119 | 0.0904 | 0.09 | 0.12 | 0.11028 |  |  | 0.1172
| 1 | 5 | 0.12 | 0.1241 | 0.13719 | 0.107 | 0.11 | 0.12 | 0.13 | 0.12 | 0.12 | 0.11 | 0.11733 | 0.14 | 0.119 | 0.13 | 0.123 | 0.0911 | 0.09 | 0.11 | 0.11374 |  |  | 103
| 1 | 6 | 0.12 | 0.1217 | 0.13528 | 0.107 | 0.11 | 0.12 | 0.13 | 0.12 | 0.12 | 0.11 | 0.1165 | 0.14 | 0.116 | 0.13 | 0.12 | 0.0987 | 0.09 | 0.11 | 0.10572
| Mean |  | 0.115 | 0.12246 | 0.13287 | 0.10767 | 0.11167 | 0.11833 | 0.13 | 0.11667 | 0.12 | 0.10833 | 0.11657 | 0.14167 | 0.11767 | 0.13167 | 0.12067 | 0.095817 | 0.09 | 0.11333 | 0.10873
| Median |  | 0.12 | 0.12188 | 0.13386 | 0.1075 | 0.11 | 0.12 | 0.13 | 0.12 | 0.12 | 0.11 | 0.11672 | 0.14 | 0.1175 | 0.13 | 0.1205 | 0.09685 | 0.09 | 0.11 | 0.10862
| Std Dev. |  | 0.012247 | 0.0014996 | 0.0050775 | 0.0017512 | 0.0040825 | 0.0040825 | 0 | 0.005164 | 0 | 0.0040825 | 0.0012887 | 0.0040825 | 0.0033267 | 0.0040825 | 0.0013663 | 0.0042508 | 1.5202e-17 | 0.005164 | 0.0037411
| Rel.Std.Dev. |  | 10.65% | 1.225% | 3.821% | 1.626% | 3.656% | 3.45% | 0% | 4.426% | 0% | 3.768% | 1.106% | 2.882% | 2.827% | 3.101% | 1.132% | 4.436% | 1.689e-14% | 4.556% | 3.441%
| PDM3 |  | -1.879% | 4.486% | 13.37% | -8.136% | -4.723% | 0.9648% | 10.92% | -0.4573% | 2.387% | -7.567% | -0.5403% | 20.87% | 0.3959% | 12.34% | 2.956% | -18.25% | -23.21% | -3.301% | -7.23%
| Z-Score (Absolute) |  | 0.19 | 0.67 | 1.89 | 1.05 | 0.58 | 0.19 | 1.55 | 0 | 0.39 | 0.97 | 0.01 | 2.91 | 0.12 | 1.75 | 0.47 | 2.43 | 3.1 | 0.39 | 0.92
| Analytical results for Nb in OREAS 242b (Certified Value 0.12 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 2
| 1 | 1 | 0.08 | 0.11507 | 0.2 | 1.12 | 0.12 | 0.14 | 0.1 | 0.07 | 0.18 | < 0.2 | < 0.2 | 0.43 | [9 empty] | 1
| 1 | 2 | 0.1 | 0.1075 | 0.2 | 1.11 | 0.11 | 0.14 | 0.09 | 0.08 | 0.18 | < 0.2 | < 0.2 | 0.43 | [9 empty] | 29
| 1 | 3 | 0.09 | 0.11737 | 0.2 | 1.1 | 0.12 | 0.13 | 0.11 | 0.08 | 0.18 | < 0.2 | < 0.2 | 0.44 | [9 empty] | 16
| 1 | 4 | 0.09 | 0.099068 | 0.3 | 1.07 | 0.12 | 0.14 | 0.1 | 0.09 | 0.25 | < 0.2 | < 0.2 | 0.43 | [9 empty] | 0.11751
| 1 | 5 | 0.1 | 0.11705 | 0.2 | 1.13 | 0.11 | 0.14 | 0.11 | 0.07 | 0.18 | < 0.2 | < 0.2 | 0.36 | [9 empty] | 104
| 1 | 6 | 0.1 | 0.12918 | 0.2 | 1.1 | 0.11 | 0.14 | 0.11 | 0.08 | 0.18 | < 0.2 | < 0.2 | 0.39
| Mean |  | 0.093333 | 0.11421 | 0.21667 | 1.105 | 0.115 | 0.13833 | 0.10333 | 0.078333 | 0.19167 |  |  | 0.41333
| Median |  | 0.095 | 0.11606 | 0.2 | 1.105 | 0.115 | 0.14 | 0.105 | 0.08 | 0.18 |  |  | 0.43
| Std Dev. |  | 0.008165 | 0.010169 | 0.040825 | 0.020736 | 0.0054772 | 0.0040825 | 0.008165 | 0.0075277 | 0.028577 |  |  | 0.031411
| Rel.Std.Dev. |  | 8.748% | 8.904% | 18.84% | 1.877% | 4.763% | 2.951% | 7.902% | 9.61% | 14.91% |  |  | 7.599%
| PDM3 |  | -20.57% | -2.807% | 84.39% | 840.4% | -2.132% | 17.72% | -12.06% | -33.34% | 63.11% |  |  | 251.8%
| Z-Score (Absolute) |  | 0.67 | 0 | NA | 32.01 | 0.03 | 0.78 | 0.35 | 1.16 | 2.5 | 0.46 | 0.46 | 9.66
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 242b (Indicative Value 5.83 ppm) | [22 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 5.309 | 5.3878 | 9.6546 | 6.13 | 6.24 | [16 empty] | 1
| 1 | 2 | 5.555 | 5.2666 | 9.7168 | 6.01 | 6.59 | [16 empty] | 4
| 1 | 3 | 5.286 | 5.4682 | 9.6308 | 6.37 | 6.41 | [16 empty] | 16
| 1 | 4 | 5.422 | 5.6496 | 9.6381 | 6.27 | 6.22 | [16 empty] | 5.8304
| 1 | 5 | 5.455 | 5.4134 | 9.6381 | 5.93 | 6.46 | [16 empty] | 10
| 1 | 6 | 5.44 | 5.347 | 9.7158 | 5.91 | 6.39
| Mean |  | 5.4112 | 5.4221 | 9.6657 | 6.1033 | 6.385
| Median |  | 5.431 | 5.4006 | 9.6464 | 6.07 | 6.4
| Std Dev. |  | 0.099662 | 0.13034 | 0.039958 | 0.18747 | 0.13896
| Rel.Std.Dev. |  | 1.842% | 2.404% | 0.4134% | 3.072% | 2.176%
| PDM3 |  | -7.19% | -7.003% | 65.78% | 4.681% | 9.512%
| Z-Score (Absolute) |  | 0.69 | 0.67 | 3.53 | 0 | 0.28
| Analytical results for Ni in OREAS 242b (Certified Value 48.5 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 47.3 | 49.39 | 47.656 | 45.29 | 50.2 | 46 | 49.3 | 51 | 48.9 | 48.9 | 52.6 | 47.6378 | 50 | 47 | 47 | 47.3 | 49.2 | 43.8 | 51 | 46.271 |  | 1
| 1 | 2 | 48.5 | 48.1213 | 50.344 | 46.55 | 49.1 | 45 | 47.9 | 52 | 50.1 | 47.1 | 50.7 | 46.6664 | 50 | 46.7 | 48 | 47.6 | 51.2 | 43.45 | 52 | 47.037 |  | 14
| 1 | 3 | 47.9 | 50.1657 | 49.956 | 46.04 | 49.7 | 47 | 47.1 | 51.9 | 49.1 | 49.2 | 51.3 | 46.4982 | 51 | 49.9 | 48 | 47.6 | 51.2 | 43.96 | 52 | 47.18 |  | 16
| 1 | 4 | 49.4 | 47.9659 | 50.099 | 45.42 | 49.8 | 46 | 48.9 | 52.7 | 49.3 | 48.8 | 50.5 | 45.9113 | 51 | 45 | 47 | 45.5 | 50.1 | 44.45 | 52 | 47.193 |  | 48.488
| 1 | 5 | 47.7 | 48.2814 | 49.966 | 46.52 | 49.9 | 46 | 47.5 | 50.8 | 48.9 | 49.6 | 50 | 47.6272 | 51 | 44.3 | 49 | 47.7 | 50 | 43.87 | 51 | 46.692 |  | 105
| 1 | 6 | 50.2 | 49.4535 | 50.575 | 46.43 | 48.6 | 46 | 47.1 | 53.2 | 50.5 | 46.8 | 51 | 46.1133 | 50 | 46.7 | 49 | 47.4 | 49.3 | 44.06 | 51 | 46.665
| Mean |  | 48.5 | 48.8963 | 49.766 | 46.0417 | 49.55 | 46 | 47.9667 | 51.9333 | 49.4667 | 48.4 | 51.0167 | 46.7424 | 50.5 | 46.6 | 48 | 47.1833 | 50.1667 | 43.9317 | 51.5 | 46.8397
| Median |  | 48.2 | 48.8357 | 50.0325 | 46.235 | 49.75 | 46 | 47.7 | 51.95 | 49.2 | 48.85 | 50.85 | 46.5823 | 50.5 | 46.7 | 48 | 47.5 | 50.05 | 43.915 | 51.5 | 46.8645
| Std Dev. |  | 1.1082 | 0.89553 | 1.0609 | 0.56382 | 0.58907 | 0.63246 | 0.93524 | 0.9331 | 0.67429 | 1.161 | 0.89312 | 0.7398 | 0.54772 | 1.9453 | 0.89443 | 0.83766 | 0.87788 | 0.32835 | 0.54772 | 0.36202
| Rel.Std.Dev. |  | 2.285% | 1.831% | 2.132% | 1.225% | 1.189% | 1.375% | 1.95% | 1.797% | 1.363% | 2.399% | 1.751% | 1.583% | 1.085% | 4.174% | 1.863% | 1.775% | 1.75% | 0.7474% | 1.064% | 0.7729%
| PDM3 |  | 0.02478% | 0.8421% | 2.636% | -5.045% | 2.19% | -5.131% | -1.075% | 7.106% | 2.018% | -0.1815% | 5.215% | -3.6% | 4.15% | -3.894% | -1.006% | -2.691% | 3.462% | -9.397% | 6.212% | -3.399%
| Z-Score (Absolute) |  | 0.02 | 0.18 | 0.53 | 0.98 | 0.45 | 1 | 0.2 | 1.42 | 0.41 | 0.02 | 1.04 | 0.69 | 0.83 | 0.75 | 0.18 | 0.51 | 0.7 | 1.84 | 1.24 | 0.65
| Analytical results for P in OREAS 242b (Certified Value 0.038 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0366 | 0.037578 | 0.037923 | 0.039 | 0.0369 | 0.037 | 0.038 | 0.036 | 0.038 | 0.038 | 0.035042 | 0.0404 | 0.037 | 0.0356 | 0.036 | 0.0277 | 0.04 | 0.042 | 0.038107 |  |  | 1
| 1 | 2 | 0.038 | 0.035164 | 0.040395 | 0.039 | 0.0371 | 0.036 | 0.038 | 0.036 | 0.039 | 0.038 | 0.034935 | 0.0401 | 0.035 | 0.0362 | 0.037 | 0.0306 | 0.03 | 0.04 | 0.038438 |  |  | #N/A
| 1 | 3 | 0.0376 | 0.037849 | 0.039104 | 0.038 | 0.0375 | 0.037 | 0.038 | 0.036 | 0.04 | 0.038 | 0.034714 | 0.0407 | 0.036 | 0.0364 | 0.037 | 0.0261 | 0.04 | 0.04 | 0.038175 |  |  | 16
| 1 | 4 | 0.0379 | 0.036074 | 0.04026 | 0.039 | 0.0367 | 0.038 | 0.038 | 0.036 | 0.039 | 0.038 | 0.034676 | 0.0393 | 0.035 | 0.0361 | 0.036 | 0.0278 | 0.04 | 0.042 | 0.038562 |  |  | 0.037803
| 1 | 5 | 0.0375 | 0.036082 | 0.040363 | 0.038 | 0.0371 | 0.037 | 0.038 | 0.036 | 0.039 | 0.038 | 0.035198 | 0.0396 | 0.036 | 0.0368 | 0.036 | 0.0291 | 0.04 | 0.04 | 0.037901 |  |  | 106
| 1 | 6 | 0.0378 | 0.037919 | 0.040431 | 0.038 | 0.0371 | 0.037 | 0.038 | 0.036 | 0.039 | 0.039 | 0.034942 | 0.0399 | 0.035 | 0.0369 | 0.037 | 0.0304 | 0.04 | 0.04 | 0.037945
| Mean |  | 0.037567 | 0.036778 | 0.039746 | 0.0385 | 0.037067 | 0.037 | 0.038 | 0.036 | 0.039 | 0.038167 | 0.034918 | 0.04 | 0.035667 | 0.036333 | 0.0365 | 0.028617 | 0.038333 | 0.040667 | 0.038188
| Median |  | 0.0377 | 0.03683 | 0.040312 | 0.0385 | 0.0371 | 0.037 | 0.038 | 0.036 | 0.039 | 0.038 | 0.034939 | 0.04 | 0.0355 | 0.0363 | 0.0365 | 0.02845 | 0.04 | 0.04 | 0.038141
| Std Dev. |  | 0.00050859 | 0.0011553 | 0.0010266 | 0.00054772 | 0.00026583 | 0.00063246 | 0 | 0 | 0.00063246 | 0.00040825 | 0.0001973 | 0.00051381 | 0.0008165 | 0.00048028 | 0.00054772 | 0.0017429 | 0.0040825 | 0.0010328 | 0.0002648
| Rel.Std.Dev. |  | 1.354% | 3.141% | 2.583% | 1.423% | 0.7172% | 1.709% | 0% | 0% | 1.622% | 1.07% | 0.565% | 1.285% | 2.289% | 1.322% | 1.501% | 6.09% | 10.65% | 2.54% | 0.6934%
| PDM3 |  | -0.6262% | -2.714% | 5.139% | 1.843% | -1.949% | -2.125% | 0.5201% | -4.77% | 3.165% | 0.961% | -7.633% | 5.811% | -5.652% | -3.889% | -3.448% | -24.3% | 1.402% | 7.574% | 1.018%
| Z-Score (Absolute) |  | 0 | 0.5 | 1.38 | 0.59 | 0.32 | 0.36 | 0.27 | 0.99 | 0.91 | 0.38 | 1.67 | 1.54 | 1.2 | 0.78 | 0.67 | 5.66 | 0.48 | 1.96 | 0.39
| Analytical results for Pb in OREAS 242b (Certified Value 16.9 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 16.4 | 17.8026 | 18.1125 | 17.7 | 16.2 | 16.7 | 17.2 | 16.1 | 16 | 17 | 16.8153 | 19 | 18.1 | 16.77 | 17 | 16.89 | 24.7 | 16.9 | 18 | 17.639 |  | 1
| 1 | 2 | 17 | 17.2623 | 18.0653 | 17.2 | 15.7 | 16.7 | 16.6 | 16.2 | 16.6 | 16 | 16.2147 | 21 | 18.4 | 16.01 | 18 | 16.29 | 25 | 19.7 | 17 | 16.617 |  | 16
| 1 | 3 | 16.6 | 17.5886 | 18.2525 | 17.1 | 16.1 | 16.2 | 17 | 16.4 | 16.8 | 17.7 | 16.2137 | 18 | 18.2 | 16.93 | 18 | 16.38 | 25.2 | 19.8 | 18 | 18.509 |  | 16
| 1 | 4 | 17.2 | 17.2511 | 18.1432 | 17.1 | 15.9 | 16.5 | 17 | 16 | 16.6 | 16.9 | 16.5955 | 17 | 18.4 | 15.19 | 17 | 15.87 | 25.4 | 18.6 | 17 | 17.35 |  | 16.863
| 1 | 5 | 16.9 | 17.1489 | 18.1915 | 17 | 16.2 | 16.2 | 16.8 | 16.4 | 16.8 | 17 | 16.6533 | 18 | 18 | 15.12 | 18 | 16.36 | 24.2 | 19.6 | 18 | 16.929 |  | 107
| 1 | 6 | 17.1 | 17.0267 | 18.2004 | 16.6 | 15.5 | 16.3 | 16.6 | 16.4 | 16.8 | 16.9 | 16.256 | 24 | 18.2 | 16.04 | 17 | 15.96 | 24.3 | 19.3 | 17 | 17.1
| Mean |  | 16.8667 | 17.3467 | 18.1609 | 17.1167 | 15.9333 | 16.4333 | 16.8667 | 16.25 | 16.6 | 16.9167 | 16.4581 | 19.5 | 18.2167 | 16.01 | 17.5 | 16.2917 | 24.8 | 18.9833 | 17.5 | 17.3573
| Median |  | 16.95 | 17.2567 | 18.1673 | 17.1 | 16 | 16.4 | 16.9 | 16.3 | 16.7 | 16.95 | 16.4257 | 18.5 | 18.2 | 16.025 | 17.5 | 16.325 | 24.85 | 19.45 | 17.5 | 17.225
| Std Dev. |  | 0.30768 | 0.29124 | 0.06734 | 0.35449 | 0.28752 | 0.23381 | 0.24221 | 0.17607 | 0.30984 | 0.54191 | 0.26244 | 2.5884 | 0.16021 | 0.76018 | 0.54772 | 0.36285 | 0.4858 | 1.1089 | 0.54772 | 0.66385
| Rel.Std.Dev. |  | 1.824% | 1.679% | 0.3708% | 2.071% | 1.805% | 1.423% | 1.436% | 1.083% | 1.866% | 3.203% | 1.595% | 13.27% | 0.8795% | 4.748% | 3.13% | 2.227% | 1.959% | 5.841% | 3.13% | 3.825%
| PDM3 |  | 0.02187% | 2.869% | 7.697% | 1.504% | -5.513% | -2.548% | 0.02187% | -3.635% | -1.56% | 0.3184% | -2.401% | 15.64% | 8.028% | -5.058% | 3.778% | -3.388% | 47.07% | 12.57% | 3.778% | 2.932%
| Z-Score (Absolute) |  | 0 | 0.66 | 1.78 | 0.34 | 1.28 | 0.6 | 0 | 0.85 | 0.37 | 0.07 | 0.56 | NA | 1.86 | 1.18 | NA | 0.79 | 10.9 | 2.91 | NA | 0.67
Results from laboratories 1.15, 1.18 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 242b (Indicative Value 15.5 ppb) | [22 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.14 | [19 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [19 empty] | 1
|  |  | 01g | 01g | [19 empty] | 1
| 1 | 1 | 16 | 15.9655 | [19 empty] | 1
| 1 | 2 | 17 | 14.0142 | [19 empty] | 5
| 1 | 3 | 17 | 13.868 | [19 empty] | 16
| 1 | 4 | 17 | 15.3344 | [19 empty] | 15.465
| 1 | 5 | 15 | 14.9826 | [19 empty] | 11
| 1 | 6 | 16 | 13.4158
| Mean |  | 16.3333 | 14.5967
| Median |  | 16.5 | 14.4984
| Std Dev. |  | 0.8165 | 0.98303
| Rel.Std.Dev. |  | 4.999% | 6.735%
| PDM3 |  | 5.615% | -5.615%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 242b (Certified Value 1.41 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 1.181 | 1.1763 | 1.907 | 1.4 | 1.31 | [16 empty] | 1
| 1 | 2 | 1.206 | 1.1644 | 1.9275 | 1.4 | 1.35 | [16 empty] | 33
| 1 | 3 | 1.148 | 1.2099 | 2.0225 | 1.4 | 1.32 | [16 empty] | 16
| 1 | 4 | 1.216 | 1.2609 | 1.9915 | 1.4 | 1.29 | [16 empty] | 1.4079
| 1 | 5 | 1.178 | 1.1819 | 1.9112 | 1.4 | 1.32 | [16 empty] | 108
| 1 | 6 | 1.225 | 1.1712 | 1.9175 | 1.4 | 1.33
| Mean |  | 1.1923 | 1.1941 | 1.9462 | 1.4 | 1.32
| Median |  | 1.1935 | 1.1791 | 1.9225 | 1.4 | 1.32
| Std Dev. |  | 0.028682 | 0.036274 | 0.048613 | 2.4324e-16 | 0.02
| Rel.Std.Dev. |  | 2.406% | 3.038% | 2.498% | 1.737e-14% | 1.515%
| PDM3 |  | -15.31% | -15.18% | 38.24% | -0.558% | -6.24%
| Z-Score (Absolute) |  | 0.68 | 0.67 | 3.35 | 0.43 | 0
| Analytical results for Pt in OREAS 242b (Indicative Value 12.3 ppb) | [22 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.14 | 1.21 | 1.28 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [17 empty] | 1
|  |  | 01g | 01g | 0.1g | 40g | [17 empty] | 1
| 1 | 1 | 11 | 14.3618 | 17 | < 10 | [17 empty] | 1
| 1 | 2 | 13 | 12.0787 | 19 | < 10 | [17 empty] | 1
| 1 | 3 | 12 | 11.4903 | 21 | < 10 | [17 empty] | 16
| 1 | 4 | 12 | 13.7372 | 19 | < 10 | [17 empty] | 12.3343
| 1 | 5 | 11 | 12.3777 | 19 | 10 | [17 empty] | 12
| 1 | 6 | 13 | 14.9768 | 15 | < 10
| Mean |  | 12 | 13.1704 | 18.3333 | 10
| Median |  | 12 | 13.0575 | 19 | 10
| Std Dev. |  | 0.89443 | 1.389 | 2.0656
| Rel.Std.Dev. |  | 7.454% | 10.55% | 11.27%
| PDM3 |  | -2.71% | 6.779% | 48.64% | -18.93%
| Z-Score (Absolute) |  | 0.12 | 0.12 | 1.22 | 1.44
| Analytical results for Rb in OREAS 242b (Certified Value 6.03 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 3
| 1 | 1 | 5.7 | 6.2625 | 5.4 | 6.5 | 6 | 6.2 | 5.9 | 5.9 | 6 | 5.7935 | 7.27 | 8.3 | 7.4 | [8 empty] | 1
| 1 | 2 | 6.1 | 6.2423 | 5.2 | 5.7 | 5.9 | 6.3 | 5.9 | 5.9 | 6.2 | 5.6198 | 7.22 | 8.4 | 6.5 | [8 empty] | 34
| 1 | 3 | 6.02 | 6.3669 | 5.2 | 5.6 | 5.8 | 6.3 | 5.9 | 6 | 6.4 | 5.5728 | 7.11 | 8.3 | 6.8 | [8 empty] | 16
| 1 | 4 | 6.03 | 6.5834 | 5.3 | 5.8 | 6 | 6.5 | 5.8 | 6 | 6.2 | 5.687 | 6.95 | 8.4 | 7.3 | [8 empty] | 6.0256
| 1 | 5 | 5.89 | 6.1408 | 5.2 | 6 | 5.7 | 6.2 | 5.9 | 6.1 | 6.3 | 5.7884 | 7.25 | 7.6 | 6.8 | [8 empty] | 109
| 1 | 6 | 6.04 | 6.4815 | 5.2 | 5.7 | 5.6 | 6.3 | 6 | 6 | 6.4 | 5.6357 | 7.17 | 8.1 | 7.5
| Mean |  | 5.9633 | 6.3462 | 5.25 | 5.8833 | 5.8333 | 6.3 | 5.9 | 5.9833 | 6.25 | 5.6829 | 7.1617 | 8.1833 | 7.05
| Median |  | 6.025 | 6.3147 | 5.2 | 5.75 | 5.85 | 6.3 | 5.9 | 6 | 6.25 | 5.6614 | 7.195 | 8.3 | 7.05
| Std Dev. |  | 0.14624 | 0.16431 | 0.083666 | 0.33116 | 0.1633 | 0.10954 | 0.063246 | 0.075277 | 0.15166 | 0.091332 | 0.11873 | 0.30605 | 0.40373
| Rel.Std.Dev. |  | 2.452% | 2.589% | 1.594% | 5.629% | 2.799% | 1.739% | 1.072% | 1.258% | 2.427% | 1.607% | 1.658% | 3.74% | 5.727%
| PDM3 |  | -1.034% | 5.321% | -12.87% | -2.361% | -3.191% | 4.554% | -2.085% | -0.7017% | 3.724% | -5.688% | 18.85% | 35.81% | 17%
| Z-Score (Absolute) |  | 0.04 | 0.81 | 1.65 | 0.22 | 0.34 | 0.71 | 0.19 | 0 | 0.6 | 0.67 | 2.64 | 4.94 | 2.39
| Analytical results for Re in OREAS 242b (Certified Value 0.002 ppm) | [22 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 3
| 1 | 1 |  | < 0.01 | 0.001 | < 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.001 | [9 empty] | 1
| 1 | 2 | 0.001 | < 0.01 | 0.001 | < 0.001 | 0.001 | 0.001 | 0.002 | 0.001 | 0.002 | < 0.05 | < 0.05 | 0.001 | [9 empty] | 35
| 1 | 3 | 0.001 | < 0.01 | 0.001 | 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.001 | < 0.05 | < 0.05 | 0.001 | [9 empty] | 16
| 1 | 4 | 0.001 | < 0.01 | 0.001 | < 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.001 | [9 empty] | 0.0015037
| 1 | 5 | 0.002 | < 0.01 | 0.001 | < 0.001 | 0.002 | 0.001 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.002 | [9 empty] | 110
| 1 | 6 | 0.001 | < 0.01 | < 0.001 | < 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.05 | 0.001
| Mean |  | 0.0012 |  | 0.001 | 0.001 | 0.0018333 | 0.0016667 | 0.002 | 0.0018333 | 0.0018333 |  |  | 0.0011667
| Median |  | 0.001 |  | 0.001 | 0.001 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 |  |  | 0.001
| Std Dev. |  | 0.00044721 |  | 0 |  | 0.00040825 | 0.0005164 | 0 | 0.00040825 | 0.00040825 |  |  | 0.00040825
| Rel.Std.Dev. |  | 37.27% |  | 0% |  | 22.27% | 30.98% | 0% | 22.27% | 22.27% |  |  | 34.99%
| PDM3 |  | -20.2% |  | -33.5% | -33.5% | 21.92% | 10.84% | 33% | 21.92% | 21.92% |  |  | -22.41%
| Z-Score (Absolute) |  | 0.66 | 3.29 | 0.95 | 1.3 | 0 | 0.17 | 0.17 | 0 | 0 | 24.03 | 24.03 | 0.69
| Analytical results for S in OREAS 242b (Certified Value 0.261 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 3
| 1 | 1 | 0.259 | 0.26228 | 0.24835 | 0.25001 | 0.249 | 0.25 | 0.25 | 0.28 | 0.26 | 0.29 | 0.27 | 0.27924 | 0.2619 | 0.28 | 0.25 | 0.2502 | 0.26 | 0.213 | 0.21 | 0.265 | 0.2855 | 1
| 1 | 2 | 0.268 | 0.26394 | 0.26563 | 0.25674 | 0.248 | 0.25 | 0.25 | 0.28 | 0.26 | 0.29 | 0.26 | 0.28095 | 0.2577 | 0.27 | 0.24 | 0.2417 | 0.26 | 0.209 | 0.21 | 0.26 | 0.2916 | 20
| 1 | 3 | 0.2651 | 0.26988 | 0.25539 | 0.25438 | 0.248 | 0.25 | 0.25 | 0.28 | 0.26 | 0.29 | 0.26 | 0.26911 | 0.2597 | 0.28 | 0.24 | 0.2578 | 0.26 | 0.217 | 0.19 | 0.265 | 0.2874 | 16
| 1 | 4 | 0.2694 | 0.26686 | 0.26131 | 0.25717 | 0.249 | 0.25 | 0.26 | 0.29 | 0.26 | 0.3 | 0.26 | 0.28308 | 0.2564 | 0.28 | 0.23 | 0.2532 | 0.25 | 0.204 | 0.21 | 0.27 | 0.2936 | 0.26064
| 1 | 5 | 0.2663 | 0.26666 | 0.2597 | 0.25348 | 0.245 | 0.25 | 0.25 | 0.28 | 0.26 | 0.29 | 0.26 | 0.27545 | 0.256 | 0.28 | 0.24 | 0.2652 | 0.26 | 0.211 | 0.2 | 0.27 | 0.2887 | 111
| 1 | 6 | 0.2676 | 0.26643 | 0.26188 | 0.25647 | 0.246 | 0.25 | 0.25 | 0.28 | 0.26 | 0.29 | 0.27 | 0.26985 | 0.2626 | 0.27 | 0.24 | 0.2629 | 0.26 | 0.214 | 0.21 | 0.265 | 0.2826
| Mean |  | 0.2659 | 0.26601 | 0.25871 | 0.25471 | 0.2475 | 0.25 | 0.25167 | 0.28167 | 0.26 | 0.29167 | 0.26333 | 0.27628 | 0.25905 | 0.27667 | 0.24 | 0.25517 | 0.25833 | 0.21133 | 0.205 | 0.26583 | 0.28823
| Median |  | 0.26695 | 0.26654 | 0.26051 | 0.25543 | 0.248 | 0.25 | 0.25 | 0.28 | 0.26 | 0.29 | 0.26 | 0.27735 | 0.2587 | 0.28 | 0.24 | 0.2555 | 0.26 | 0.212 | 0.21 | 0.265 | 0.28805
| Std Dev. |  | 0.0036867 | 0.0026269 | 0.0060675 | 0.0027201 | 0.0016432 | 0 | 0.0040825 | 0.0040825 | 0 | 0.0040825 | 0.005164 | 0.0058361 | 0.0028034 | 0.005164 | 0.0063246 | 0.0086867 | 0.0040825 | 0.0045019 | 0.0083666 | 0.0037639 | 0.0040063
| Rel.Std.Dev. |  | 1.387% | 0.9875% | 2.345% | 1.068% | 0.6639% | 0% | 1.622% | 1.449% | 0% | 1.4% | 1.961% | 2.112% | 1.082% | 1.866% | 2.635% | 3.404% | 1.58% | 2.13% | 4.081% | 1.416% | 1.39%
| PDM3 |  | 2.019% | 2.06% | -0.74% | -2.274% | -5.04% | -4.081% | -3.442% | 8.069% | -0.2443% | 11.91% | 1.035% | 6.002% | -0.6088% | 6.15% | -7.918% | -2.099% | -0.8837% | -18.92% | -21.35% | 1.994% | 10.59%
| Z-Score (Absolute) |  | 0.63 | 0.64 | 0.03 | 0.4 | 1.05 | 0.83 | 0.67 | 2.07 | 0.09 | 2.98 | 0.39 | 1.57 | 0 | 1.61 | 1.74 | 0.35 | 0.07 | 4.36 | 4.94 | 0.62 | 2.67
| Analytical results for Sb in OREAS 242b (Certified Value 1.67 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 2
| 1 | 1 | 1.97 | 1.5234 | 1.6236 | 1.84 | 1.35 | 1.86 | 2.02 | 1.67 | 1.7 | 2.01 | 1.0213 | 1.14 | 1.24 | 1.82 | 1.57 | 1.73 | < 5 | 1.7 |  |  |  | 1
| 1 | 2 | 2.06 | 1.5235 | 1.5512 | 1.76 | 1.23 | 1.83 | 2.08 | 1.68 | 1.76 | 1.95 | 0.98516 | 1.18 | 1.29 | 1.79 | 1.6 | 1.81 | < 5 | 1.6 |  |  |  | 36
| 1 | 3 | 2.02 | 1.4891 | 1.4479 | 1.76 | 1.18 | 1.85 | 2.03 | 1.72 | 1.86 | 2.01 | 1.0282 | 1.19 | 1.27 | 1.88 | 1.53 | 1.77 | < 5 | 1.8 |  |  |  | 16
| 1 | 4 | 2.11 | 1.5903 | 1.5497 | 1.62 | 1.34 | 1.89 | 2.19 | 1.67 | 1.88 | 2 | 0.98376 | 1.09 | 1.25 | 1.81 | 1.52 | 1.76 | < 5 | 1.6 |  |  |  | 1.6737
| 1 | 5 | 2.05 | 1.5013 | 1.5799 | 1.76 | 1.34 | 1.77 | 2.02 | 1.74 | 1.82 | 2 | 1.0364 | 1.14 | 1.19 | 1.73 | 1.61 | 1.7 | < 5 | 1.8 |  |  |  | 112
| 1 | 6 | 2.17 | 1.5229 | 1.8062 | 1.64 | 1.34 | 1.82 | 2.06 | 1.73 | 1.86 | 2.07 | 1.0246 | 1.13 | 1.27 | 1.86 | 1.53 | 1.72 | < 5 | 1.6
| Mean |  | 2.0633 | 1.5251 | 1.5931 | 1.73 | 1.2967 | 1.8367 | 2.0667 | 1.7017 | 1.8133 | 2.0067 | 1.0132 | 1.145 | 1.2517 | 1.815 | 1.56 | 1.7483 |  | 1.6833
| Median |  | 2.055 | 1.5232 | 1.5656 | 1.76 | 1.34 | 1.84 | 2.045 | 1.7 | 1.84 | 2.005 | 1.023 | 1.14 | 1.26 | 1.815 | 1.55 | 1.745 |  | 1.65
| Std Dev. |  | 0.069761 | 0.034994 | 0.11937 | 0.083666 | 0.072847 | 0.040825 | 0.065013 | 0.031885 | 0.070048 | 0.038297 | 0.022856 | 0.036194 | 0.034881 | 0.053198 | 0.038987 | 0.039707 |  | 0.098319
| Rel.Std.Dev. |  | 3.381% | 2.295% | 7.493% | 4.836% | 5.618% | 2.223% | 3.146% | 1.874% | 3.863% | 1.908% | 2.256% | 3.161% | 2.787% | 2.931% | 2.499% | 2.271% |  | 5.841%
| PDM3 |  | 23.28% | -8.88% | -4.817% | 3.363% | -22.53% | 9.736% | 23.48% | 1.67% | 8.341% | 19.89% | -39.46% | -31.59% | -25.22% | 8.441% | -6.794% | 4.458% |  | 0.5743%
| Z-Score (Absolute) |  | 1.35 | 0.74 | 0.48 | 0.06 | 1.63 | 0.47 | 1.37 | 0.06 | 0.38 | 1.13 | 2.73 | 2.22 | 1.81 | 0.39 | 0.61 | 0.13 | 3.05 | 0.13
| Analytical results for Sc in OREAS 242b (Certified Value 4.82 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 2
| 1 | 1 | 4.3 | 4.7689 | 6.7 | 5.7 | 4.9 | 5.6 | 5.1 | 4.5 | 5.4 | 5.0128 | 6 | 4.6 | 6 | 3.7 | 5 | 4 | 3 | 3.795 |  |  |  | 1
| 1 | 2 | 4.5 | 4.8111 | 6.2 | 5.5 | 4.8 | 5.6 | 5.2 | 4.5 | 5.3 | 4.7367 | 6 | 4.2 | 7 | 3.8 | 5 | 4.2 | 3 | 3.79 |  |  |  | 37
| 1 | 3 | 4.4 | 4.9709 | 6.3 | 5.3 | 4.7 | 5.5 | 5.1 | 4.7 | 5 | 5.1476 | 6 | 4.7 | 6 | 3.8 | 5 | 4.2 | 3 | 3.947 |  |  |  | 16
| 1 | 4 | 4.5 | 4.6612 | 6.2 | 5.5 | 5 | 5.8 | 5 | 4.8 | 5.1 | 5.1716 | 6 | 4.1 | 6 | 3.5 | 5 | 4.3 | 3 | 4.141 |  |  |  | 4.8192
| 1 | 5 | 4.5 | 4.747 | 6.1 | 5.5 | 4.6 | 5.5 | 5.1 | 4.5 | 5.1 | 5.3253 | 6 | 4.3 | 7 | 3.8 | 4 | 4.2 | 3 | 3.655 |  |  |  | 113
| 1 | 6 | 4.6 | 5.068 | 6.1 | 5.1 | 4.7 | 5.7 | 5.1 | 4.6 | 5.3 | 4.8179 | 6 | 4.6 | 7 | 3.8 | 5 | 4.2 | 3 | 3.688
| Mean |  | 4.4667 | 4.8378 | 6.2667 | 5.4333 | 4.7833 | 5.6167 | 5.1 | 4.6 | 5.2 | 5.0353 | 6 | 4.4167 | 6.5 | 3.7333 | 4.8333 | 4.1833 | 3 | 3.836
| Median |  | 4.5 | 4.79 | 6.2 | 5.5 | 4.75 | 5.6 | 5.1 | 4.55 | 5.2 | 5.0802 | 6 | 4.45 | 6.5 | 3.8 | 5 | 4.2 | 3 | 3.7925
| Std Dev. |  | 0.10328 | 0.15206 | 0.22509 | 0.20656 | 0.1472 | 0.1169 | 0.063246 | 0.12649 | 0.15492 | 0.22459 | 0 | 0.24833 | 0.54772 | 0.12111 | 0.40825 | 0.098319 | 0 | 0.181
| Rel.Std.Dev. |  | 2.312% | 3.143% | 3.592% | 3.802% | 3.077% | 2.081% | 1.24% | 2.75% | 2.979% | 4.46% | 0% | 5.623% | 8.427% | 3.244% | 8.447% | 2.35% | 0% | 4.718%
| PDM3 |  | -7.316% | 0.3861% | 30.03% | 12.74% | -0.7448% | 16.55% | 5.826% | -4.549% | 7.901% | 4.484% | 24.5% | -8.353% | 34.88% | -22.53% | 0.2927% | -13.19% | -37.75% | -20.4%
| Z-Score (Absolute) |  | 0.59 | 0.05 | 2.51 | 1.07 | 0.05 | 1.39 | 0.5 | 0.36 | 0.67 | 0.39 | NA | 0.68 | NA | 1.85 | NA | 1.08 | NA | 1.68
Results from laboratories 1.16, 1.18, 1.21 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 242b (Certified Value 0.88 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1.5 | 1.1487 | 0.7 | < 5 | 0.8 | 0.9 | 0.9 | 0.9 | 1.1 | < 1 | 0.75 | < 1 | 1.1 | 0.9 | 1.64 | 1 | [5 empty] | 1
| 1 | 2 | < 1.5 | 1.1016 | 0.9 | < 5 | 0.8 | 1 | 0.8 | 0.9 | 1 | < 1 | 0.77 | < 1 | 1.1 | 0.8 | 1.61 | 1 | 0.585 |  |  |  |  | 2
| 1 | 3 | < 1.5 | 1.2297 | 0.8 | < 5 | 0.9 | 0.7 | 0.8 | 0.8 | 1.3 | < 1 | 0.78 | < 1 | 0.8 | 0.9 | 1.63 | 1 | 0.655 |  |  |  |  | 16
| 1 | 4 | < 1.5 | 1.135 | 0.8 | < 5 | 0.9 | 1 | 0.8 | 1.1 | 0.9 | < 1 | 0.76 | < 1 | 0.6 | 0.9 | 1.47 | 1 | 2.259 |  |  |  |  | 0.87503
| 1 | 5 | < 1.5 | 1.1729 | 0.8 | < 5 | 0.9 | 1 | 0.9 | 0.9 | 0.9 | < 1 | 0.79 | < 1 | 0.5 | 0.9 | 1.2 | 1 | [5 empty] | 114
| 1 | 6 | < 1.5 | 1.0678 | 0.7 | < 5 | 1 | 0.8 | 0.9 | 1.2 | 1.2 | < 1 | 0.73 | < 1 | 0.7 | 0.9 | 1.57 | 1 | 0.518
| Mean |  |  | 1.1426 | 0.78333 |  | 0.88333 | 0.9 | 0.85 | 0.96667 | 1.0667 |  | 0.76333 |  | 0.8 | 0.88333 | 1.52 | 1 | 1.0042
| Median |  |  | 1.1419 | 0.8 |  | 0.9 | 0.95 | 0.85 | 0.9 | 1.05 |  | 0.765 |  | 0.75 | 0.9 | 1.59 | 1 | 0.62
| Std Dev. |  |  | 0.056361 | 0.075277 |  | 0.075277 | 0.12649 | 0.054772 | 0.15055 | 0.1633 |  | 0.021602 |  | 0.25298 | 0.040825 | 0.16852 | 0 | 0.83837
| Rel.Std.Dev. |  |  | 4.933% | 9.61% |  | 8.522% | 14.05% | 6.444% | 15.57% | 15.31% |  | 2.83% |  | 31.62% | 4.622% | 11.09% | 0% | 83.48%
| PDM3 |  |  | 30.58% | -10.48% |  | 0.9493% | 2.854% | -2.86% | 10.47% | 21.9% |  | -12.76% |  | -8.574% | 0.9493% | 73.71% | 14.28% | 14.77%
| Z-Score (Absolute) |  | 0.75 | 1.46 | 0.56 | NA | 0 | 0.09 | 0.19 | 0.47 | 1.03 | 2.15 | 0.67 | 2.15 | 0.47 | 0 | 3.58 | NA | 0.68
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.08 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 242b (Indicative Value 1.44 ppm) | [22 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 1.453 | 1.4573 | 2.4424 | 1.9 | 1.36 | [16 empty] | 1
| 1 | 2 | 1.519 | 1.4157 | 2.4872 | 1.8 | 1.42 | [16 empty] | 7
| 1 | 3 | 1.398 | 1.4608 | 2.4586 | 1.5 | 1.46 | [16 empty] | 16
| 1 | 4 | 1.499 | 1.4628 | 2.4999 | 1.7 | 1.38 | [16 empty] | 1.4414
| 1 | 5 | 1.497 | 1.4256 | 2.4906 | 1.8 | 1.44 | [16 empty] | 13
| 1 | 6 | 1.465 | 1.4528 | 2.5134 | 1.9 | 1.38
| Mean |  | 1.4718 | 1.4458 | 2.482 | 1.7667 | 1.4067
| Median |  | 1.481 | 1.455 | 2.4889 | 1.8 | 1.4
| Std Dev. |  | 0.043471 | 0.020049 | 0.026539 | 0.15055 | 0.039328
| Rel.Std.Dev. |  | 2.954% | 1.387% | 1.069% | 8.522% | 2.796%
| PDM3 |  | 2.108% | 0.3041% | 72.19% | 22.56% | -2.413%
| Z-Score (Absolute) |  | 0 | 0.27 | 10.45 | 3.05 | 0.67
| Analytical results for Sn in OREAS 242b (Certified Value 0.63 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 2
| 1 | 1 | 0.47 | 0.5391 | 0.15663 | 0.69 | 0.8 | 0.6 | 0.8 | 0.5 | 0.6 | 0.5 | 0.6052 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 1
| 1 | 2 | 0.55 | 0.58159 | 0.15344 | 0.65 | 0.8 | 0.6 | 0.8 | 0.5 | 0.5 | 0.6 | 0.53771 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 11
| 1 | 3 | 0.52 | 0.51759 | 0.10805 | 0.65 | 0.8 | 0.6 | 0.7 | 0.5 | 0.6 | 0.6 | 0.56443 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 16
| 1 | 4 | 0.56 | 0.59293 | 0.11499 | 0.64 | 0.7 | 0.6 | 0.8 | 0.5 | 0.6 | 0.6 | 0.61749 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 0.62646
| 1 | 5 | 0.57 | 0.54757 | 0.1891 | 0.66 | 0.8 | 0.5 | 0.8 | 0.5 | 0.6 | 0.6 | 0.57101 | 0.7 | 0.7 | < 5 | < 1 | [6 empty] | 115
| 1 | 6 | 0.59 | 0.59007 | 0.2655 | 0.63 | 0.7 | 0.5 | 0.8 | 0.5 | 0.6 | 0.6 | 0.5601 | 0.7 | 0.7 | < 5 | < 1
| Mean |  | 0.54333 | 0.56147 | 0.16462 | 0.65333 | 0.76667 | 0.56667 | 0.78333 | 0.5 | 0.58333 | 0.58333 | 0.57599 | 0.7 | 0.7
| Median |  | 0.555 | 0.56458 | 0.15504 | 0.65 | 0.8 | 0.6 | 0.8 | 0.5 | 0.6 | 0.6 | 0.56772 | 0.7 | 0.7
| Std Dev. |  | 0.04274 | 0.031086 | 0.057667 | 0.020656 | 0.05164 | 0.05164 | 0.040825 | 0 | 0.040825 | 0.040825 | 0.029837 | 1.2162e-16 | 1.2162e-16
| Rel.Std.Dev. |  | 7.866% | 5.537% | 35.03% | 3.162% | 6.736% | 9.113% | 5.212% | 0% | 6.999% | 6.999% | 5.18% | 1.737e-14% | 1.737e-14%
| PDM3 |  | -13.27% | -10.37% | -73.72% | 4.29% | 22.38% | -9.544% | 25.04% | -20.19% | -6.883% | -6.883% | -8.056% | 11.74% | 11.74%
| Z-Score (Absolute) |  | 0.32 | 0.18 | 3.39 | 0.57 | 1.48 | 0.13 | 1.62 | 0.67 | 0 | 0 | 0.06 | 0.94 | 0.94 | 15.51 | 0.67
| Analytical results for Sr in OREAS 242b (Certified Value 30.3 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 1
| 1 | 1 | 30.99 | 31.7401 | 20.407 | 58 | 36.6 | 32 | 33.9 | 29.4 | 26.3 | 28.9 | 32.7297 | 54.3 | 33.3 | 59 | 24.1 | 44.2 | 32.4 | 24 | 22.185 |  |  | 1
| 1 | 2 | 32.31 | 33.7432 | 21.316 | 52.1 | 35.7 | 31.7 | 33.5 | 30.5 | 26.8 | 29.5 | 32.9318 | 55.1 | 29.3 | 61 | 25.3 | 39.6 | 35 | 27 | 22.683 |  |  | 12
| 1 | 3 | 32.11 | 34.1893 | 21.79 | 53.1 | 36.2 | 30.6 | 34.2 | 30.2 | 27.3 | 30.8 | 32.6029 | 54.7 | 31.5 | 58 | 25 | 38.6 | 32.1 | 26 | 22.95 |  |  | 16
| 1 | 4 | 32.86 | 32.0098 | 21.142 | 52.2 | 34.4 | 32.3 | 34.5 | 29.4 | 26.8 | 30.2 | 33.1357 | 54.8 | 29.4 | 60 | 22.5 | 38.1 | 33.9 | 25 | 24.129 |  |  | 30.2731
| 1 | 5 | 32.31 | 32.4086 | 21.708 | 52.8 | 36.8 | 30.2 | 33.1 | 29.5 | 26.9 | 29.4 | 33.0811 | 54.1 | 33.2 | 60 | 25.3 | 52.6 | 32.1 | 27 | 21.299 |  |  | 116
| 1 | 6 | 32.81 | 33.3556 | 21.454 | 52.1 | 35.9 | 30.2 | 34.1 | 30.6 | 26.7 | 31.7 | 32.675 | 54.5 | 32.5 | 60 | 25.1 | 40.8 | 31.9 | 24 | 21.957
| Mean |  | 32.2317 | 32.9078 | 21.3028 | 53.3833 | 35.9333 | 31.1667 | 33.8833 | 29.9333 | 26.8 | 30.0833 | 32.8594 | 54.5833 | 31.5333 | 59.6667 | 24.55 | 42.3167 | 32.9 | 25.5 | 22.5338
| Median |  | 32.31 | 32.8821 | 21.385 | 52.5 | 36.05 | 31.15 | 34 | 29.85 | 26.8 | 29.85 | 32.8307 | 54.6 | 32 | 60 | 25.05 | 40.2 | 32.25 | 25.5 | 22.434
| Std Dev. |  | 0.67795 | 0.996 | 0.50058 | 2.2991 | 0.85713 | 0.94375 | 0.50761 | 0.56451 | 0.32249 | 1.0342 | 0.22243 | 0.36009 | 1.8096 | 1.0328 | 1.0986 | 5.4876 | 1.2602 | 1.3784 | 0.97152
| Rel.Std.Dev. |  | 2.103% | 3.027% | 2.35% | 4.307% | 2.385% | 3.028% | 1.498% | 1.886% | 1.203% | 3.438% | 0.6769% | 0.6597% | 5.739% | 1.731% | 4.475% | 12.97% | 3.83% | 5.406% | 4.311%
| PDM3 |  | 6.47% | 8.703% | -29.63% | 76.34% | 18.7% | 2.952% | 11.93% | -1.122% | -11.47% | -0.6267% | 8.543% | 80.3% | 4.163% | 97.09% | -18.9% | 39.78% | 8.678% | -15.77% | -25.56%
| Z-Score (Absolute) |  | 0 | 0.12 | 1.99 | 3.85 | 0.67 | 0.19 | 0.3 | 0.42 | 0.99 | 0.39 | 0.11 | 4.07 | 0.13 | 5 | 1.4 | 1.84 | 0.12 | 1.23 | 1.77
| Analytical results for Ta in OREAS 242b (Certified Value < 0.01 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 25
| 1 | 3 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 117
| 1 | 6 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01
| Mean |  |  |  |  | 0.01 | 0.01 | 0.01
| Median |  |  |  |  | 0.01 | 0.01 | 0.01
| Std Dev. | [5 empty] | 0
| Rel.Std.Dev. | [5 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 15.51 | 0.67 | 0 | 2.02 | 0 | 0.67 | 0.67 | 15.51 | 15.51 | 0.67
| Analytical results for Tb in OREAS 242b (Certified Value 0.31 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.261 | 0.30443 | 0.35704 | 0.3 | 0.29 | [16 empty] | 1
| 1 | 2 | 0.281 | 0.29431 | 0.38512 | 0.3 | 0.305 | [16 empty] | 26
| 1 | 3 | 0.271 | 0.30866 | 0.31676 | 0.3 | 0.315 | [16 empty] | 16
| 1 | 4 | 0.274 | 0.32004 | 0.35428 | 0.3 | 0.3 | [16 empty] | 0.30998
| 1 | 5 | 0.286 | 0.30683 | 0.39196 | 0.3 | 0.305 | [16 empty] | 118
| 1 | 6 | 0.272 | 0.30878 | 0.39628 | 0.3 | 0.295
| Mean |  | 0.27417 | 0.30718 | 0.36691 | 0.3 | 0.30167
| Median |  | 0.273 | 0.30775 | 0.37108 | 0.3 | 0.3025
| Std Dev. |  | 0.0086583 | 0.0082862 | 0.030311 | 0 | 0.008756
| Rel.Std.Dev. |  | 3.158% | 2.698% | 8.261% | 0% | 2.903%
| PDM3 |  | -11.55% | -0.9058% | 18.36% | -3.221% | -2.683%
| Z-Score (Absolute) |  | 3.37 | 0.67 | 7.99 | 0.2 | 0
| Analytical results for Te in OREAS 242b (Certified Value 0.15 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | [5 empty] | 2
| 1 | 1 | 0.1 | < 0.2 | 0.15 | 0.35 | 0.14 | 0.18 | 0.12 | 0.15 | 0.18 | 0.12014 | 0.15 | 0.14 | 0.08 | 0.15 | 0.15 | < 0.2 | [5 empty] | 1
| 1 | 2 | 0.2 | < 0.2 | 0.2 | 0.36 | 0.16 | 0.18 | 0.16 | 0.15 | 0.14 | 0.12481 | 0.14 | 0.13 | 0.1 | 0.14 | 0.15 | < 0.2 | [5 empty] | 27
| 1 | 3 | 0.2 | < 0.2 | 0.16 | 0.33 | 0.17 | 0.2 | 0.14 | 0.14 | 0.14 | 0.14742 | 0.13 | 0.14 | 0.18 | 0.15 | 0.15 | < 0.2 | [5 empty] | 16
| 1 | 4 | 0.2 | < 0.2 | 0.15 | 0.34 | 0.15 | 0.15 | 0.14 | 0.14 | 0.15 | 0.15296 | 0.11 | 0.15 | 0.18 | 0.14 | 0.15 | < 0.2 | [5 empty] | 0.14621
| 1 | 5 | 0.2 | < 0.2 | 0.15 | 0.35 | 0.15 | 0.14 | 0.15 | 0.13 | 0.15 | 0.12296 | 0.12 | 0.15 | 0.17 | 0.14 | 0.14 | < 0.2 | [5 empty] | 119
| 1 | 6 | 0.2 | < 0.2 | 0.15 | 0.34 | 0.14 | 0.16 | 0.15 | 0.16 | 0.15 | 0.15675 | 0.13 | 0.14 | 0.09 | 0.15 | 0.16 | < 0.2
| Mean |  | 0.18333 |  | 0.16 | 0.345 | 0.15167 | 0.16833 | 0.14333 | 0.145 | 0.15167 | 0.13751 | 0.13 | 0.14167 | 0.13333 | 0.145 | 0.15
| Median |  | 0.2 |  | 0.15 | 0.345 | 0.15 | 0.17 | 0.145 | 0.145 | 0.15 | 0.13611 | 0.13 | 0.14 | 0.135 | 0.145 | 0.15
| Std Dev. |  | 0.040825 |  | 0.02 | 0.010488 | 0.01169 | 0.022286 | 0.013663 | 0.010488 | 0.01472 | 0.016625 | 0.014142 | 0.0075277 | 0.048028 | 0.0054772 | 0.0063246
| Rel.Std.Dev. |  | 22.27% |  | 12.5% | 3.04% | 7.708% | 13.24% | 9.532% | 7.233% | 9.705% | 12.09% | 10.88% | 5.314% | 36.02% | 3.777% | 4.216%
| PDM3 |  | 25.39% |  | 9.432% | 136% | 3.733% | 15.13% | -1.967% | -0.8269% | 3.733% | -5.951% | -11.09% | -3.107% | -8.806% | -0.8269% | 2.593%
| Z-Score (Absolute) |  | 2.7 | 3.17 | 1.06 | 14.08 | 0.47 | 1.64 | 0.12 | 0 | 0.47 | 0.53 | 1.06 | 0.23 | 0.82 | 0 | 0.35 | 3.17
| Analytical results for Th in OREAS 242b (Certified Value 0.92 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 2
| 1 | 1 | 0.84 | 0.89086 | 3.858 | 1.1724 | 1 | 1.1 | 0.9 | 1 | 0.8 | 0.8 | 0.9 | 0.92666 | 1.01 | 0.8 | 0.8 | 0.96 | < 5 | 1 |  |  |  | 1
| 1 | 2 | 0.86 | 0.82674 | 4.102 | 1.126 | 1 | 1.1 | 0.9 | 0.9 | 0.8 | 0.8 | 0.9 | 0.90764 | 1.03 | 0.8 | 0.8 | 0.93 | < 5 | 1.05 |  |  |  | 28
| 1 | 3 | 0.82 | 0.89022 | 3.936 | 1.1086 | 1 | 1.2 | 0.9 | 0.9 | 0.8 | 0.9 | 1 | 0.87395 | 0.99 | 0.9 | 0.8 | 0.94 | < 5 | 0.95 |  |  |  | 16
| 1 | 4 | 0.86 | 0.87484 | 4.133 | 1.167 | 1 | 1.1 | 0.9 | 0.9 | 0.8 | 0.8 | 0.9 | 0.91198 | 1 | 0.8 | 0.8 | 0.96 | < 5 | 1 |  |  |  | 0.92249
| 1 | 5 | 0.85 | 0.84614 | 4.221 | 1.1313 | 1 | 1.2 | 0.8 | 0.9 | 0.8 | 0.9 | 0.9 | 0.94114 | 1.02 | 0.8 | 0.8 | 0.86 | < 5 | 1.1 |  |  |  | 120
| 1 | 6 | 0.89 | 0.8646 | 3.788 | 1.1423 | 1 | 1.1 | 0.9 | 0.9 | 0.8 | 0.8 | 0.9 | 0.87733 | 1.01 | 0.8 | 0.8 | 0.93 | < 5 | 1
| Mean |  | 0.85333 | 0.86557 | 4.0063 | 1.1413 | 1 | 1.1333 | 0.88333 | 0.91667 | 0.8 | 0.83333 | 0.91667 | 0.90645 | 1.01 | 0.81667 | 0.8 | 0.93 |  | 1.0167
| Median |  | 0.855 | 0.86972 | 4.019 | 1.1368 | 1 | 1.1 | 0.9 | 0.9 | 0.8 | 0.8 | 0.9 | 0.90981 | 1.01 | 0.8 | 0.8 | 0.935 |  | 1
| Std Dev. |  | 0.023381 | 0.025377 | 0.17082 | 0.024653 | 0 | 0.05164 | 0.040825 | 0.040825 | 1.2162e-16 | 0.05164 | 0.040825 | 0.026638 | 0.014142 | 0.040825 | 1.2162e-16 | 0.036878 |  | 0.05164
| Rel.Std.Dev. |  | 2.74% | 2.932% | 4.264% | 2.16% | 0% | 4.556% | 4.622% | 4.454% | 1.52e-14% | 6.197% | 4.454% | 2.939% | 1.4% | 4.999% | 1.52e-14% | 3.965% |  | 5.079%
| PDM3 |  | -7.496% | -6.17% | 334.3% | 23.72% | 8.403% | 22.86% | -4.244% | -0.6307% | -13.28% | -9.664% | -0.6307% | -1.738% | 9.487% | -11.47% | -13.28% | 0.8146% |  | 10.21%
| Z-Score (Absolute) |  | 0.46 | 0.37 | 22.32 | 1.62 | NA | 1.57 | 0.24 | 0 | 0.84 | 0.6 | 0 | 0.07 | 0.67 | 0.72 | 0.84 | 0.1 | 11.44 | 0.72
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 242b (Certified Value 0.319 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.2842 | 0.3129 | 0.14653 | 0.39 | 0.3 | 0.302 | 0.355 | 0.283 | 0.291 | 0.287 | 0.40398 | 0.43 | 0.37 | 0.3888 | 0.217 | 0.289 | 0.32 | 0.12 | 0.28031 |  |  | 1
| 1 | 2 | 0.3068 | 0.32273 | 0.15447 | 0.39 | 0.29 | 0.304 | 0.349 | 0.289 | 0.296 | 0.298 | 0.38946 | 0.423 | 0.319 | 0.4053 | 0.239 | 0.281 | 0.32 | 0.115 | 0.27898 |  |  | 29
| 1 | 3 | 0.2978 | 0.33268 | 0.16095 | 0.39 | 0.3 | 0.316 | 0.351 | 0.286 | 0.308 | 0.291 | 0.37797 | 0.435 | 0.35 | 0.4045 | 0.23 | 0.277 | 0.32 | 0.115 | 0.29065 |  |  | 16
| 1 | 4 | 0.3053 | 0.31165 | 0.15093 | 0.38 | 0.3 | 0.318 | 0.355 | 0.29 | 0.299 | 0.302 | 0.41847 | 0.422 | 0.316 | 0.401 | 0.199 | 0.259 | 0.33 | 0.12 | 0.3098 |  |  | 0.31904
| 1 | 5 | 0.308 | 0.31375 | 0.16181 | 0.38 | 0.29 | 0.313 | 0.351 | 0.292 | 0.299 | 0.297 | 0.40513 | 0.419 | 0.356 | 0.4087 | 0.245 | 0.262 | 0.32 | 0.125 | 0.26367 |  |  | 121
| 1 | 6 | 0.3115 | 0.32949 | 0.15161 | 0.38 | 0.3 | 0.301 | 0.346 | 0.287 | 0.298 | 0.306 | 0.38206 | 0.429 | 0.358 | 0.4112 | 0.236 | 0.283 | 0.32 | 0.125 | 0.26629
| Mean |  | 0.30227 | 0.32053 | 0.15438 | 0.385 | 0.29667 | 0.309 | 0.35117 | 0.28783 | 0.2985 | 0.29683 | 0.39618 | 0.42633 | 0.34483 | 0.40325 | 0.22767 | 0.27517 | 0.32167 | 0.12 | 0.28162
| Median |  | 0.30605 | 0.31824 | 0.15304 | 0.385 | 0.3 | 0.3085 | 0.351 | 0.288 | 0.2985 | 0.2975 | 0.39672 | 0.426 | 0.353 | 0.4049 | 0.233 | 0.279 | 0.32 | 0.12 | 0.27965
| Std Dev. |  | 0.0099422 | 0.0091191 | 0.0059923 | 0.0054772 | 0.005164 | 0.0075366 | 0.0034881 | 0.0031885 | 0.0055408 | 0.006969 | 0.01558 | 0.0059889 | 0.022167 | 0.0079061 | 0.016967 | 0.01204 | 0.0040825 | 0.0044721 | 0.016977
| Rel.Std.Dev. |  | 3.289% | 2.845% | 3.881% | 1.423% | 1.741% | 2.439% | 0.9933% | 1.108% | 1.856% | 2.348% | 3.933% | 1.405% | 6.428% | 1.961% | 7.452% | 4.376% | 1.269% | 3.727% | 6.028%
| PDM3 |  | -5.258% | 0.4675% | -51.61% | 20.67% | -7.013% | -3.148% | 10.07% | -9.782% | -6.439% | -6.961% | 24.18% | 33.63% | 8.084% | 26.39% | -28.64% | -13.75% | 0.8225% | -62.39% | -11.73%
| Z-Score (Absolute) |  | 0 | 0.45 | 3.68 | 2.06 | 0.14 | 0.17 | 1.22 | 0.36 | 0.09 | 0.14 | 2.34 | 3.09 | 1.06 | 2.51 | 1.86 | 0.67 | 0.48 | 4.54 | 0.51
| Analytical results for Tl in OREAS 242b (Certified Value 0.094 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 3
| 1 | 1 | 0.09 | 0.090062 | 0.079165 | 0.1 | 0.32 | 0.09 | 0.09 | 0.09 | 0.08 | 0.1 | 0.095499 | 0.092 | 0.11 | 0.08 | 0.1 | 0.12 | < 5 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.09 | 0.087749 | 0.082945 | 0.1 | 0.29 | 0.09 | 0.1 | 0.09 | 0.08 | 0.1 | 0.099707 | 0.095 | 0.11 | 0.08 | 0.1 | 0.12 | < 5 | < 0.1 |  |  |  | 30
| 1 | 3 | 0.1 | 0.091581 | 0.09026 | 0.1 | 0.31 | 0.09 | 0.1 | 0.09 | 0.09 | 0.1 | 0.1026 | 0.098 | 0.1 | 0.09 | 0.1 | 0.12 | < 5 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.1 | 0.09892 | 0.080658 | 0.1 | 0.3 | 0.09 | 0.1 | 0.09 | 0.09 | 0.1 | 0.09576 | 0.091 | 0.1 | 0.08 | 0.09 | 0.12 | < 5 | < 0.1 |  |  |  | 0.093992
| 1 | 5 | 0.1 | 0.088549 | 0.090548 | 0.1 | 0.31 | 0.09 | 0.1 | 0.1 | 0.09 | 0.1 | 0.10308 | 0.093 | 0.11 | 0.09 | 0.1 | 0.12 | < 5 | < 0.1 |  |  |  | 122
| 1 | 6 | 0.1 | 0.09172 | 0.081704 | 0.09 | 0.3 | 0.1 | 0.1 | 0.09 | 0.09 | 0.09 | 0.098825 | 0.096 | 0.1 | 0.08 | 0.1 | 0.12 | < 5 | < 0.1
| Mean |  | 0.096667 | 0.09143 | 0.084213 | 0.098333 | 0.305 | 0.091667 | 0.098333 | 0.091667 | 0.086667 | 0.098333 | 0.099245 | 0.094167 | 0.105 | 0.083333 | 0.098333 | 0.12
| Median |  | 0.1 | 0.090821 | 0.082324 | 0.1 | 0.305 | 0.09 | 0.1 | 0.09 | 0.09 | 0.1 | 0.099266 | 0.094 | 0.105 | 0.08 | 0.1 | 0.12
| Std Dev. |  | 0.005164 | 0.0039984 | 0.0049542 | 0.0040825 | 0.010488 | 0.0040825 | 0.0040825 | 0.0040825 | 0.005164 | 0.0040825 | 0.0032406 | 0.0026394 | 0.0054772 | 0.005164 | 0.0040825 | 0
| Rel.Std.Dev. |  | 5.342% | 4.373% | 5.883% | 4.152% | 3.439% | 4.454% | 4.152% | 4.454% | 5.958% | 4.152% | 3.265% | 2.803% | 5.216% | 6.197% | 4.152% | 0%
| PDM3 |  | 2.845% | -2.726% | -10.4% | 4.619% | 224.5% | -2.474% | 4.619% | -2.474% | -7.794% | 4.619% | 5.588% | 0.1857% | 11.71% | -11.34% | 4.619% | 27.67%
| Z-Score (Absolute) |  | 0.09 | 0.69 | 1.51 | 0.09 | 23.51 | 0.66 | 0.09 | 0.66 | 1.23 | 0.09 | 0.2 | 0.38 | 0.85 | 1.61 | 0.09 | 2.55 | 272.2 | 5.38
| Analytical results for Tm in OREAS 242b (Indicative Value 0.17 ppm) | [22 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 0.162 | 0.15871 | 0.2 | 0.15 | [17 empty] | 1
| 1 | 2 | 0.166 | 0.15662 | 0.2 | 0.135 | [17 empty] | 4
| 1 | 3 | 0.153 | 0.16226 | 0.2 | 0.14 | [17 empty] | 16
| 1 | 4 | 0.166 | 0.16769 | 0.2 | 0.155 | [17 empty] | 0.1679
| 1 | 5 | 0.171 | 0.16729 | 0.2 | 0.155 | [17 empty] | 10
| 1 | 6 | 0.174 | 0.15991 | 0.2 | 0.13
| Mean |  | 0.16533 | 0.16208 | 0.2 | 0.14417
| Median |  | 0.166 | 0.16108 | 0.2 | 0.145
| Std Dev. |  | 0.0073666 | 0.0045733 | 3.0405e-17 | 0.010685
| Rel.Std.Dev. |  | 4.456% | 2.822% | 1.52e-14% | 7.411%
| PDM3 |  | -1.526% | -3.463% | 19.12% | -14.13%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 2.31 | 1.24
| Analytical results for U in OREAS 242b (Certified Value 0.24 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g |  |  |  | 3
| 1 | 1 | 0.2 | 0.26921 | 0.24201 | 0.25372 | 0.3 | 0.23 | 0.24 | 0.25 | 0.23 | 0.23 | 0.24 | 0.27309 | 0.27 | 0.28 | 0.23 | 0.2 | 0.24 | 0.2 |  |  |  | 1
| 1 | 2 | 0.25 | 0.22712 | 0.24106 | 0.27004 | 0.3 | 0.25 | 0.25 | 0.26 | 0.23 | 0.24 | 0.24 | 0.26752 | 0.27 | 0.29 | 0.23 | 0.2 | 0.23 | 0.25 |  |  |  | 32
| 1 | 3 | 0.23 | 0.23905 | 0.24879 | 0.23744 | 0.3 | 0.25 | 0.24 | 0.25 | 0.24 | 0.24 | 0.27 | 0.24955 | 0.28 | 0.27 | 0.26 | 0.2 | 0.24 | 0.25 |  |  |  | 16
| 1 | 4 | 0.24 | 0.23419 | 0.24729 | 0.23164 | 0.3 | 0.24 | 0.24 | 0.26 | 0.24 | 0.24 | 0.25 | 0.26169 | 0.26 | 0.27 | 0.22 | 0.2 | 0.25 | 0.2 |  |  |  | 0.24356
| 1 | 5 | 0.24 | 0.23458 | 0.24508 | 0.2566 | 0.3 | 0.25 | 0.23 | 0.26 | 0.23 | 0.25 | 0.23 | 0.2662 | 0.28 | 0.29 | 0.22 | 0.2 | 0.22 | 0.2 |  |  |  | 123
| 1 | 6 | 0.28 | 0.23586 | 0.24915 | 0.2818 | 0.3 | 0.24 | 0.26 | 0.25 | 0.23 | 0.24 | 0.23 | 0.25307 | 0.27 | 0.28 | 0.25 | 0.2 | 0.23 | 0.2
| Mean |  | 0.24 | 0.24 | 0.24556 | 0.25521 | 0.3 | 0.24333 | 0.24333 | 0.255 | 0.23333 | 0.24 | 0.24333 | 0.26185 | 0.27167 | 0.28 | 0.235 | 0.2 | 0.235 | 0.21667
| Median |  | 0.24 | 0.23522 | 0.24618 | 0.25516 | 0.3 | 0.245 | 0.24 | 0.255 | 0.23 | 0.24 | 0.24 | 0.26395 | 0.27 | 0.28 | 0.23 | 0.2 | 0.235 | 0.2
| Std Dev. |  | 0.026077 | 0.014836 | 0.0034466 | 0.018986 | 0 | 0.008165 | 0.010328 | 0.0054772 | 0.005164 | 0.0063246 | 0.015055 | 0.0090098 | 0.0075277 | 0.0089443 | 0.016432 | 3.0405e-17 | 0.010488 | 0.02582
| Rel.Std.Dev. |  | 10.87% | 6.181% | 1.404% | 7.44% | 0% | 3.355% | 4.244% | 2.148% | 2.213% | 2.635% | 6.187% | 3.441% | 2.771% | 3.194% | 6.992% | 1.52e-14% | 4.463% | 11.92%
| PDM3 |  | -1.463% | -1.462% | 0.8209% | 4.78% | 23.17% | -0.09443% | -0.09443% | 4.696% | -4.2% | -1.463% | -0.09443% | 7.509% | 11.54% | 14.96% | -3.516% | -17.89% | -3.516% | -11.04%
| Z-Score (Absolute) |  | 0.27 | 0.27 | 0.18 | 0.96 | 4.59 | 0 | 0 | 0.94 | 0.81 | 0.27 | 0 | 1.5 | 2.29 | 2.97 | 0.67 | NA | 0.67 | 2.16
Results from laboratory 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 242b (Certified Value 125 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 0
| 1 | 1 | 122 | 128.56 | 77.071 | 157 | 132 | 120 | 132 | 118 | 125 | 113 | 142.3105 | 172 | 134 | 160 | 107 | 159 | 118 | 90 | 118.277 |  |  | 1
| 1 | 2 | 129 | 131.8144 | 82.177 | 152 | 132 | 120 | 131 | 121 | 126 | 116 | 137.9713 | 170 | 119 | 164 | 112 | 167 | 118 | 90 | 118.885 |  |  | 33
| 1 | 3 | 126 | 134.0597 | 81.175 | 153 | 134 | 122 | 130 | 119 | 128 | 115 | 135.1083 | 172 | 127 | 163 | 111 | 165 | 120.6 | 90 | 122.659 |  |  | 16
| 1 | 4 | 129 | 127.02 | 82.429 | 153 | 132 | 124 | 132 | 119 | 126 | 117 | 147.0247 | 174 | 120 | 162 | 103 | 165 | 122 | 90 | 127.586 |  |  | 124.7271
| 1 | 5 | 130 | 128.41 | 84.671 | 152 | 132 | 121 | 131 | 120 | 127 | 115 | 148.1302 | 171 | 130 | 167 | 114 | 165 | 119.2 | 90 | 112.596 |  |  | 124
| 1 | 6 | 130 | 134.8791 | 82.921 | 153 | 132 | 119 | 131 | 120 | 125 | 119 | 135.9177 | 171 | 128 | 167 | 113 | 162 | 118.1 | 90 | 114.805
| Mean |  | 127.6667 | 130.7905 | 81.7407 | 153.3333 | 132.3333 | 121 | 131.1667 | 119.5 | 126.1667 | 115.8333 | 141.0771 | 171.6667 | 126.3333 | 163.8333 | 110 | 163.8333 | 119.3167 | 90 | 119.1347
| Median |  | 129 | 130.1872 | 82.303 | 153 | 132 | 120.5 | 131 | 119.5 | 126 | 115.5 | 140.1409 | 171.5 | 127.5 | 163.5 | 111.5 | 165 | 118.65 | 90 | 118.581
| Std Dev. |  | 3.1411 | 3.2653 | 2.5601 | 1.8619 | 0.8165 | 1.7889 | 0.75277 | 1.0488 | 1.169 | 2.0412 | 5.6309 | 1.3663 | 5.8195 | 2.7869 | 4.1952 | 2.8577 | 1.6618 | 0 | 5.4064
| Rel.Std.Dev. |  | 2.46% | 2.497% | 3.132% | 1.214% | 0.617% | 1.478% | 0.5739% | 0.8777% | 0.9266% | 1.762% | 3.991% | 0.7959% | 4.606% | 1.701% | 3.814% | 1.744% | 1.393% | 0% | 4.538%
| PDM3 |  | 2.357% | 4.861% | -34.46% | 22.94% | 6.098% | -2.988% | 5.163% | -4.191% | 1.154% | -7.131% | 13.11% | 37.63% | 1.288% | 31.35% | -11.81% | 31.35% | -4.338% | -27.84% | -4.484%
| Z-Score (Absolute) |  | 0.12 | 0.42 | 4.18 | 2.53 | 0.56 | 0.5 | 0.45 | 0.64 | 0.02 | 0.98 | 1.38 | 4.25 | 0 | 3.51 | 1.53 | 3.51 | 0.66 | 3.4 | 0.67
| Analytical results for W in OREAS 242b (Certified Value 10.6 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g |  |  |  |  | 1
| 1 | 1 | 8.99 | 11.1284 | 15.932 | 11 | 12.01 | 11.95 | 13.35 | 10.75 | 10.4 | 11.85 | 10.7463 | 8.2 | 8.84 | 8 | 9.1 | 12.7 | < 5 |  |  |  |  | 1
| 1 | 2 | 8.87 | 11.0835 | 16.753 | 10.3 | 11.91 | 12.1 | 12.85 | 10.65 | 10.85 | 11.3 | 10.281 | 8 | 9 | 8 | 9.2 | 12.8 | < 5 |  |  |  |  | 34
| 1 | 3 | 9.02 | 11.1006 | 17.509 | 10.4 | 11.64 | 12.15 | 13.1 | 10.95 | 11.05 | 11.8 | 10.2311 | 8.1 | 9.28 | 8 | 8.7 | 12.9 | < 5 |  |  |  |  | 16
| 1 | 4 | 8.81 | 11.7611 | 16.722 | 9.5 | 11.16 | 12 | 13.5 | 10.95 | 11.05 | 11.55 | 10.0224 | 7.7 | 8.84 | 8 | 8.5 | 12.9 | < 5 |  |  |  |  | 10.6455
| 1 | 5 | 8.67 | 11.2688 | 15.857 | 10.5 | 12.46 | 11.65 | 13.05 | 11.25 | 10.8 | 11.45 | 10.5373 | 7.6 | 8.11 | 8 | 9.3 | 12.4 | < 5 |  |  |  |  | 125
| 1 | 6 | 8.47 | 11.1227 | 16.576 | 9.6 | 11.45 | 11.75 | 12.7 | 11 | 10.6 | 11.45 | 10.6452 | 8.6 | 8.53 | 8 | 8.6 | 12.4 | < 5
| Mean |  | 8.805 | 11.2442 | 16.5582 | 10.2167 | 11.7717 | 11.9333 | 13.0917 | 10.925 | 10.7917 | 11.5667 | 10.4106 | 8.0333 | 8.7667 | 8 | 8.9 | 12.6833
| Median |  | 8.84 | 11.1255 | 16.649 | 10.35 | 11.775 | 11.975 | 13.075 | 10.95 | 10.825 | 11.5 | 10.4092 | 8.05 | 8.84 | 8 | 8.9 | 12.75
| Std Dev. |  | 0.20744 | 0.26171 | 0.60876 | 0.57067 | 0.4571 | 0.19664 | 0.29903 | 0.20917 | 0.25577 | 0.21602 | 0.27694 | 0.36148 | 0.40396 | 0 | 0.34059 | 0.23166
| Rel.Std.Dev. |  | 2.356% | 2.328% | 3.676% | 5.586% | 3.883% | 1.648% | 2.284% | 1.915% | 2.37% | 1.868% | 2.66% | 4.5% | 4.608% | 0% | 3.827% | 1.826%
| PDM3 |  | -17.29% | 5.624% | 55.54% | -4.028% | 10.58% | 12.1% | 22.98% | 2.626% | 1.373% | 8.654% | -2.207% | -24.54% | -17.65% | -24.85% | -16.4% | 19.14%
| Z-Score (Absolute) |  | 0.95 | 0.18 | 2.65 | 0.3 | 0.42 | 0.5 | 1.04 | 0.03 | 0.03 | 0.33 | 0.21 | 1.31 | 0.97 | NA | 0.91 | 0.85 | 3.89
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 242b (Certified Value 10.7 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 2
| 1 | 1 | 9.89 | 10.5188 | 9.887 | 11.9 | 11.01 | 11.15 | 11.6 | 10.4 | 10.25 | 10.25 | 11.041 | 14.1 | 13 | 10.1 | 9.561 | [6 empty] | 1
| 1 | 2 | 10.62 | 10.7298 | 9.86 | 11.4 | 10.84 | 11.1 | 12 | 10.55 | 10.85 | 10.3 | 10.6842 | 14.1 | 13.3 | 10.8 | 9.539 | [6 empty] | 35
| 1 | 3 | 10.31 | 10.8421 | 10.422 | 11.8 | 10.89 | 10.9 | 11.75 | 10.6 | 10.7 | 10.75 | 10.5735 | 14 | 13 | 10.7 | 9.712 | [6 empty] | 16
| 1 | 4 | 10.53 | 10.6361 | 9.842 | 11.7 | 10.63 | 11.2 | 12.05 | 10.6 | 10.6 | 10.6 | 10.9747 | 14.2 | 13.2 | 10.1 | 9.994 | [6 empty] | 10.6909
| 1 | 5 | 10.34 | 10.6638 | 10.139 | 11.7 | 11.13 | 10.65 | 11.7 | 10.7 | 10.65 | 10.4 | 10.8482 | 13.7 | 12.7 | 10.5 | 9.272 | [6 empty] | 126
| 1 | 6 | 10.67 | 10.9428 | 10.089 | 11.6 | 10.81 | 10.65 | 12.35 | 11.25 | 10.8 | 10.9 | 10.5182 | 14.2 | 13.2 | 10.3 | 9.368
| Mean |  | 10.3933 | 10.7223 | 10.0398 | 11.6833 | 10.885 | 10.9417 | 11.9083 | 10.6833 | 10.6417 | 10.5333 | 10.7733 | 14.05 | 13.0667 | 10.4167 | 9.5743
| Median |  | 10.435 | 10.6968 | 9.988 | 11.7 | 10.865 | 11 | 11.875 | 10.6 | 10.675 | 10.5 | 10.7662 | 14.1 | 13.1 | 10.4 | 9.55
| Std Dev. |  | 0.28626 | 0.15185 | 0.22501 | 0.17224 | 0.17225 | 0.24782 | 0.27824 | 0.29439 | 0.21311 | 0.26013 | 0.21482 | 0.18708 | 0.21602 | 0.29944 | 0.25709
| Rel.Std.Dev. |  | 2.754% | 1.416% | 2.241% | 1.474% | 1.582% | 2.265% | 2.337% | 2.756% | 2.003% | 2.47% | 1.994% | 1.332% | 1.653% | 2.875% | 2.685%
| PDM3 |  | -2.784% | 0.2929% | -6.09% | 9.283% | 1.815% | 2.345% | 11.39% | -0.07119% | -0.4609% | -1.474% | 0.7703% | 31.42% | 22.22% | -2.566% | -10.44%
| Z-Score (Absolute) |  | 0.73 | 0 | 1.51 | 2.12 | 0.36 | 0.48 | 2.62 | 0.09 | 0.18 | 0.42 | 0.11 | 7.34 | 5.17 | 0.67 | 2.53
| Analytical results for Yb in OREAS 242b (Indicative Value 1.26 ppm) | [22 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.848 | 0.97682 | 1.8586 | 1.3 | [17 empty] | 1
| 1 | 2 | 1.05 | 0.9564 | 1.8147 | 1.2 | [17 empty] | 5
| 1 | 3 | 0.961 | 1.001 | 1.8884 | 1.2 | [17 empty] | 16
| 1 | 4 | 1.002 | 0.98586 | 1.8616 | 1.2 | [17 empty] | 1.2635
| 1 | 5 | 1.021 | 0.96533 | 1.809 | 1.2 | [17 empty] | 11
| 1 | 6 | 1.039 | 1.0228 | 1.7966 | 1.2
| Mean |  | 0.98683 | 0.98471 | 1.8382 | 1.2167
| Median |  | 1.0115 | 0.98134 | 1.8367 | 1.2
| Std Dev. |  | 0.074901 | 0.024323 | 0.0364 | 0.040825
| Rel.Std.Dev. |  | 7.59% | 2.47% | 1.98% | 3.355%
| PDM3 |  | -21.9% | -22.07% | 45.48% | -3.709%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 4.28 | 0.67
| Analytical results for Zn in OREAS 242b (Certified Value 101 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 0
| 1 | 1 | 94 | 100.7745 | 99.501 | 98.993 | 103 | 105 | 99 | 102 | 98 | 106 | 103 | 101.7789 | 94 | 107 | 95.4 | 99 | 94 | 91.2 | 108.8 | 109 | 99.135 | 1
| 1 | 2 | 96 | 99.9147 | 103.031 | 100.281 | 100 | 105 | 97 | 102 | 99 | 107 | 102 | 99.8775 | 96 | 106 | 92.1 | 101 | 95.1 | 96.5 | 129.4 | 108 | 100.875 | 15
| 1 | 3 | 96 | 102.2193 | 100.171 | 99.4185 | 102 | 107 | 99 | 101 | 98 | 107 | 104 | 100.4081 | 98 | 109 | 91.7 | 102 | 96.8 | 93.9 | 111.3 | 105 | 103.903 | 16
| 1 | 4 | 98 | 97.7037 | 103.972 | 99.6945 | 101 | 106 | 100 | 103 | 97 | 107 | 102 | 101.2608 | 91 | 105 | 91.6 | 99 | 93.8 | 92.6 | 112.6 | 108 | 100.479 | 100.5041
| 1 | 5 | 95 | 98.0936 | 104.966 | 100.695 | 102 | 113 | 99 | 102 | 99 | 108 | 102 | 101.8971 | 95 | 107 | 93.9 | 104 | 98.4 | 96.6 | 131.9 | 101 | 99.706 | 127
| 1 | 6 | 98 | 101.9947 | 102.64 | 98.9585 | 99.6 | 101 | 99 | 103 | 99 | 106 | 104 | 100.268 | 93 | 108 | 90.1 | 104 | 95 | 95.3 | 105 | 110 | 102.283
| Mean |  | 96.1667 | 100.1168 | 102.3802 | 99.6734 | 101.2667 | 106.1667 | 98.8333 | 102.1667 | 98.3333 | 106.8333 | 102.8333 | 100.9151 | 94.5 | 107 | 92.4667 | 101.5 | 95.5167 | 94.35 | 116.5 | 106.8333 | 101.0635
| Median |  | 96 | 100.3446 | 102.8355 | 99.5565 | 101.5 | 105.5 | 99 | 102 | 98.5 | 107 | 102.5 | 100.8344 | 94.5 | 107 | 91.9 | 101.5 | 95.05 | 94.6 | 111.95 | 108 | 100.677
| Std Dev. |  | 1.6021 | 1.915 | 2.1392 | 0.70033 | 1.3064 | 3.92 | 0.98319 | 0.75277 | 0.8165 | 0.75277 | 0.98319 | 0.84645 | 2.429 | 1.4142 | 1.8833 | 2.2583 | 1.7691 | 2.1806 | 11.2903 | 3.3116 | 1.7601
| Rel.Std.Dev. |  | 1.666% | 1.913% | 2.089% | 0.7026% | 1.29% | 3.692% | 0.9948% | 0.7368% | 0.8303% | 0.7046% | 0.9561% | 0.8388% | 2.57% | 1.322% | 2.037% | 2.225% | 1.852% | 2.311% | 9.691% | 3.1% | 1.742%
| PDM3 |  | -4.316% | -0.3854% | 1.867% | -0.8265% | 0.7587% | 5.634% | -1.662% | 1.654% | -2.16% | 6.297% | 2.318% | 0.4089% | -5.974% | 6.463% | -7.997% | 0.9909% | -4.962% | -6.123% | 15.92% | 6.297% | 0.5566%
| Z-Score (Absolute) |  | 1.21 | 0.23 | 0.33 | 0.34 | 0.05 | 1.26 | 0.55 | 0.27 | 0.67 | 1.43 | 0.44 | 0.04 | 1.62 | 1.47 | 2.12 | 0.11 | 1.37 | 1.66 | 3.81 | 1.43 | 0
| Analytical results for Zr in OREAS 242b (Certified Value 16 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 2
| 1 | 1 | 14.1 | 15.6904 | 23.3044 | 14.5 | 17.4 | 15.6 | 18.2 | 15.5 | 15.6 | 15.6 | 17.3314 | 17.6 | 21.2 | 14 | 2 | [6 empty] | 1
| 1 | 2 | 15 | 15.6921 | 23.9806 | 14 | 16.9 | 15.5 | 18 | 15.4 | 15.8 | 16 | 16.8532 | 17 | 21.3 | 14.8 | 2 | [6 empty] | 16
| 1 | 3 | 14.6 | 16.0267 | 23.6572 | 15.2 | 16.4 | 16 | 18 | 15.8 | 15.7 | 15.7 | 16.8759 | 16.4 | 21.7 | 14.8 | 2 | [6 empty] | 16
| 1 | 4 | 14.3 | 15.5172 | 23.6768 | 11.6 | 16.5 | 15.7 | 18.7 | 15.4 | 15.8 | 16.1 | 17.1789 | 16.6 | 21.8 | 15.5 | 2 | [6 empty] | 15.9908
| 1 | 5 | 15 | 15.6101 | 23.9414 | 14 | 17.7 | 15.6 | 18.4 | 15.9 | 15.2 | 16.2 | 17.4631 | 17.5 | 20.2 | 15.9 | 2 | [6 empty] | 128
| 1 | 6 | 15.3 | 16.2927 | 23.716 | 12.6 | 17.2 | 15.1 | 18.4 | 15.9 | 15.8 | 16.5 | 16.6928 | 17.2 | 21.2 | 15 | 2
| Mean |  | 14.7167 | 15.8049 | 23.7127 | 13.65 | 17.0167 | 15.5833 | 18.2833 | 15.65 | 15.65 | 16.0167 | 17.0659 | 17.05 | 21.2333 | 15 | 2
| Median |  | 14.8 | 15.6913 | 23.6964 | 14 | 17.05 | 15.6 | 18.3 | 15.65 | 15.75 | 16.05 | 17.0274 | 17.1 | 21.25 | 14.9 | 2
| Std Dev. |  | 0.46224 | 0.29449 | 0.24297 | 1.3172 | 0.51153 | 0.29269 | 0.27142 | 0.2429 | 0.23452 | 0.33116 | 0.30384 | 0.48062 | 0.56804 | 0.65422 | 0
| Rel.Std.Dev. |  | 3.141% | 1.863% | 1.025% | 9.65% | 3.006% | 1.878% | 1.484% | 1.552% | 1.499% | 2.068% | 1.78% | 2.819% | 2.675% | 4.361% | 0%
| PDM3 |  | -7.968% | -1.163% | 48.29% | -14.64% | 6.415% | -2.548% | 14.34% | -2.131% | -2.131% | 0.1616% | 6.723% | 6.624% | 32.78% | -6.196% | -87.49%
| Z-Score (Absolute) |  | 0.72 | 0.06 | 4.69 | 1.36 | 0.66 | 0.2 | 1.43 | 0.16 | 0.16 | 0.06 | 0.69 | 0.68 | 3.2 | 0.55 | NA
Results from laboratory 1.28 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1055: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 242b (Indicative Value 12.78 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.79 | 1
| 1 | 2 | 12.76 | 9
| Mean |  | 12.775 | 16
| Median |  | 12.775 | 12.775
| Std Dev. |  | 0.021213 | 15
| Rel.Std.Dev. |  | 0.1661%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 242b (Indicative Value 60 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 60 | 10
| Mean |  | 60 | 16
| Median |  | 60 | 60
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 242b (Indicative Value 190 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 179 | 1
| 1 | 2 | 201 | 11
| Mean |  | 190 | 16
| Median |  | 190 | 189.805
| Std Dev. |  | 15.5563 | 17
| Rel.Std.Dev. |  | 8.188%
| PDM3 |  | 0.1027%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 242b (Indicative Value 9.19 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 9.18 | 1
| 1 | 2 | 9.19 | 12
| Mean |  | 9.185 | 16
| Median |  | 9.185 | 9.185
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.07698%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 242b (Indicative Value 1250 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1250 | 1
| 1 | 2 | 1250 | 13
| Mean |  | 1250 | 16
| Median |  | 1250 | 1250
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 242b (Indicative Value 50 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 14
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 242b (Indicative Value 212 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 205 | 1
| 1 | 2 | 219 | 15
| Mean |  | 212 | 16
| Median |  | 212 | 211.9175
| Std Dev. |  | 9.8995 | 21
| Rel.Std.Dev. |  | 4.67%
| PDM3 |  | 0.03893%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 242b (Indicative Value 205 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 210 | 1
| 1 | 2 | 200 | 16
| Mean |  | 205 | 16
| Median |  | 205 | 205
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 242b (Indicative Value 10.96 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.952 | 1
| 1 | 2 | 10.966 | 9
| Mean |  | 10.959 | 16
| Median |  | 10.959 | 10.9587
| Std Dev. |  | 0.0098995 | 15
| Rel.Std.Dev. |  | 0.09033%
| PDM3 |  | 0.003189%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 242b (Indicative Value 0.447 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.448 | 1
| 1 | 2 | 0.446 | 10
| Mean |  | 0.447 | 16
| Median |  | 0.447 | 0.447
| Std Dev. |  | 0.0014142 | 16
| Rel.Std.Dev. |  | 0.3164%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 242b (Indicative Value 5.93 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.93 | 1
| 1 | 2 | 5.93 | 11
| Mean |  | 5.93 | 16
| Median |  | 5.93 | 5.93
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 242b (Indicative Value 0.171 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.171 | 1
| 1 | 2 | 0.171 | 12
| Mean |  | 0.171 | 16
| Median |  | 0.171 | 0.171
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 242b (Indicative Value 2.33 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.33 | 1
| 1 | 2 | 2.33 | 13
| Mean |  | 2.33 | 16
| Median |  | 2.33 | 2.33
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 242b (Indicative Value 130 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 150 | 1
| 1 | 2 | 110 | 14
| Mean |  | 130 | 16
| Median |  | 130 | 130
| Std Dev. |  | 28.2843 | 20
| Rel.Std.Dev. |  | 21.76%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 242b (Indicative Value 0.088 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0894 | 1
| 1 | 2 | 0.0871 | 15
| Mean |  | 0.08825 | 16
| Median |  | 0.08825 | 0.088227
| Std Dev. |  | 0.0016263 | 21
| Rel.Std.Dev. |  | 1.843%
| PDM3 |  | 0.02652%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 242b (Indicative Value 85 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 90 | 16
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 242b (Indicative Value 0.25 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.25 | 20
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 242b (Indicative Value 53.8 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 53.81 | 1
| 1 | 2 | 53.78 | 10
| Mean |  | 53.795 | 16
| Median |  | 53.795 | 53.795
| Std Dev. |  | 0.021213 | 16
| Rel.Std.Dev. |  | 0.03943%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 242b (Indicative Value 40 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 40 | 11
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 242b (Indicative Value 200 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 200 | 1
| 1 | 2 | 200 | 12
| Mean |  | 200 | 16
| Median |  | 200 | 200
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 242b (Indicative Value 1.02 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.014 | 1
| 1 | 2 | 1.018 | 13
| Mean |  | 1.016 | 16
| Median |  | 1.016 | 1.016
| Std Dev. |  | 0.0028284 | 19
| Rel.Std.Dev. |  | 0.2784%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 242b (Indicative Value 482 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 482 | 1
| 1 | 2 | 482 | 14
| Mean |  | 482 | 16
| Median |  | 482 | 482.004
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.0008299%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 242b (Indicative Value 125 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 120 | 1
| 1 | 2 | 130 | 15
| Mean |  | 125 | 16
| Median |  | 125 | 125
| Std Dev. |  | 7.0711 | 21
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 242b (Indicative Value 100 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 16
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 242b (Indicative Value 2.97 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.94 | 1
| 1 | 2 | 2.99 | 18
| Mean |  | 2.965 | 16
| Median |  | 2.965 | 2.965
| Std Dev. |  | 0.035355 | 24
| Rel.Std.Dev. |  | 1.192%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 242b (Indicative Value 0.14 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.14 | 1
| 1 | 2 | 0.14 | 20
| Mean |  | 0.14 | 16
| Median |  | 0.14 | 0.14
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 242b (Indicative Value 0.27 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.28 | 1
| 1 | 2 | 0.26 | 20
| Mean |  | 0.27 | 16
| Median |  | 0.27 | 0.27
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 242b (Indicative Value 4.1 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.1 | 1
| 1 | 2 | 4.1 | 22
| Mean |  | 4.1 | 16
| Median |  | 4.1 | 4.1
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 242b (Indicative Value 53 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 52.2 | 1
| 1 | 2 | 52.8 | 10
| Mean |  | 52.5 | 16
| Median |  | 52.5 | 52.5
| Std Dev. |  | 0.42426 | 29
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 242b (Indicative Value 155 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 157 | 1
| 1 | 2 | 152 | 24
| Mean |  | 154.5 | 16
| Median |  | 154.5 | 154.5
| Std Dev. |  | 3.5355 | 30
| Rel.Std.Dev. |  | 2.288%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 242b (Indicative Value 0.5 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.4 | 25
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0.14142 | 31
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 242b (Indicative Value 0.4 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 26
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 242b (Indicative Value 0.45 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.4 | 27
| Mean |  | 0.45 | 16
| Median |  | 0.45 | 0.45
| Std Dev. |  | 0.070711 | 33
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 242b (Indicative Value 16.9 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.2 | 1
| 1 | 2 | 16.6 | 28
| Mean |  | 16.9 | 16
| Median |  | 16.9 | 16.9
| Std Dev. |  | 0.42426 | 34
| Rel.Std.Dev. |  | 2.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 242b (Indicative Value 41.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.9 | 1
| 1 | 2 | 41.2 | 14
| Mean |  | 41.05 | 16
| Median |  | 41.05 | 41.05
| Std Dev. |  | 0.21213 | 35
| Rel.Std.Dev. |  | 0.5168%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 242b (Indicative Value 129 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 129 | 1
| 1 | 2 | 128 | 30
| Mean |  | 128.5 | 16
| Median |  | 128.5 | 128.5
| Std Dev. |  | 0.70711 | 36
| Rel.Std.Dev. |  | 0.5503%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 242b (Indicative Value 1.1 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.11 | 1
| 1 | 2 | 1.08 | 31
| Mean |  | 1.095 | 16
| Median |  | 1.095 | 1.095
| Std Dev. |  | 0.021213 | 37
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 242b (Indicative Value 165 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 162 | 1
| 1 | 2 | 168 | 16
| Mean |  | 165 | 16
| Median |  | 165 | 165
| Std Dev. |  | 4.2426 | 38
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 242b (Indicative Value 3.88 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.9 | 1
| 1 | 2 | 3.86 | 4
| Mean |  | 3.88 | 16
| Median |  | 3.88 | 3.88
| Std Dev. |  | 0.028284 | 39
| Rel.Std.Dev. |  | 0.729%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 242b (Indicative Value 2.34 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.33 | 1
| 1 | 2 | 2.34 | 5
| Mean |  | 2.335 | 16
| Median |  | 2.335 | 2.335
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.3028%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 242b (Indicative Value 0.83 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.83 | 1
| 1 | 2 | 0.83 | 35
| Mean |  | 0.83 | 16
| Median |  | 0.83 | 0.83
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 242b (Indicative Value 15.4 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.4 | 1
| 1 | 2 | 15.3 | 36
| Mean |  | 15.35 | 16
| Median |  | 15.35 | 15.35
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 0.4607%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 242b (Indicative Value 3.36 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.36 | 1
| 1 | 2 | 3.35 | 6
| Mean |  | 3.355 | 16
| Median |  | 3.355 | 3.355
| Std Dev. |  | 0.0070711 | 43
| Rel.Std.Dev. |  | 0.2108%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 242b (Indicative Value 1.3 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.3 | 38
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 242b (Indicative Value 2.12 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.11 | 1
| 1 | 2 | 2.13 | 22
| Mean |  | 2.12 | 16
| Median |  | 2.12 | 2.12
| Std Dev. |  | 0.014142 | 28
| Rel.Std.Dev. |  | 0.6671%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 242b (Indicative Value 0.84 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.83 | 1
| 1 | 2 | 0.84 | 23
| Mean |  | 0.835 | 16
| Median |  | 0.835 | 0.835
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.8468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 242b (Indicative Value 0.075 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.1 | 24
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.035355 | 30
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 242b (Indicative Value 7.76 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.72 | 1
| 1 | 2 | 7.79 | 25
| Mean |  | 7.755 | 16
| Median |  | 7.755 | 7.755
| Std Dev. |  | 0.049497 | 31
| Rel.Std.Dev. |  | 0.6383%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 242b (Indicative Value 0.33 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.33 | 1
| 1 | 2 | 0.33 | 7
| Mean |  | 0.33 | 16
| Median |  | 0.33 | 0.33
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 242b (Indicative Value 0.131 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.131 | 1
| 1 | 2 | 0.131 | 27
| Mean |  | 0.131 | 16
| Median |  | 0.131 | 0.131
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 242b (Indicative Value 8.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.6 | 1
| 1 | 2 | 8.6 | 28
| Mean |  | 8.6 | 16
| Median |  | 8.6 | 8.6
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 242b (Indicative Value 3.98 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.98 | 1
| 1 | 2 | 3.97 | 29
| Mean |  | 3.975 | 16
| Median |  | 3.975 | 3.975
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.1779%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 242b (Indicative Value 10.1 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.1 | 1
| 1 | 2 | 10.1 | 4
| Mean |  | 10.1 | 16
| Median |  | 10.1 | 10.1
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 242b (Indicative Value 83 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 88 | 1
| 1 | 2 | 78 | 14
| Mean |  | 83 | 16
| Median |  | 83 | 83
| Std Dev. |  | 7.0711 | 37
| Rel.Std.Dev. |  | 8.519%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 242b (Indicative Value 19 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19 | 1
| 1 | 2 | 19 | 16
| Mean |  | 19 | 16
| Median |  | 19 | 19
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 242b (Indicative Value 2.22 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.24 | 1
| 1 | 2 | 2.19 | 33
| Mean |  | 2.215 | 16
| Median |  | 2.215 | 2.215
| Std Dev. |  | 0.035355 | 39
| Rel.Std.Dev. |  | 1.596%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 242b (Indicative Value 14.1 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.4 | 1
| 1 | 2 | 13.7 | 34
| Mean |  | 14.05 | 16
| Median |  | 14.05 | 14.05
| Std Dev. |  | 0.49497 | 40
| Rel.Std.Dev. |  | 3.523%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 242b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 35
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 242b (Indicative Value 3.5 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3.5 | 36
| Mean |  | 3.5 | 16
| Median |  | 3.5 | 3.5
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 242b (Indicative Value 36.9 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.6 | 1
| 1 | 2 | 37.2 | 37
| Mean |  | 36.9 | 16
| Median |  | 36.9 | 36.9
| Std Dev. |  | 0.42426 | 43
| Rel.Std.Dev. |  | 1.15%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 242b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 2
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 242b (Indicative Value 2.86 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.92 | 7
| Mean |  | 2.86 | 16
| Median |  | 2.86 | 2.86
| Std Dev. |  | 0.084853 | 28
| Rel.Std.Dev. |  | 2.967%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 242b (Indicative Value 0.9 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 0.8 | 11
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 242b (Indicative Value 184 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 185 | 1
| 1 | 2 | 182 | 12
| Mean |  | 183.5 | 16
| Median |  | 183.5 | 183.5
| Std Dev. |  | 2.1213 | 30
| Rel.Std.Dev. |  | 1.156%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 242b (Indicative Value 0.28 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.27 | 25
| Mean |  | 0.28 | 16
| Median |  | 0.28 | 0.28
| Std Dev. |  | 0.014142 | 31
| Rel.Std.Dev. |  | 5.051%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 242b (Indicative Value 0.58 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.57 | 26
| Mean |  | 0.58 | 16
| Median |  | 0.58 | 0.58
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 2.438%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 242b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 242b (Indicative Value 1.63 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.61 | 1
| 1 | 2 | 1.65 | 28
| Mean |  | 1.63 | 16
| Median |  | 1.63 | 1.63
| Std Dev. |  | 0.028284 | 34
| Rel.Std.Dev. |  | 1.735%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 242b (Indicative Value 0.597 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.598 | 1
| 1 | 2 | 0.596 | 29
| Mean |  | 0.597 | 16
| Median |  | 0.597 | 0.597
| Std Dev. |  | 0.0014142 | 35
| Rel.Std.Dev. |  | 0.2369%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 242b (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 30
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 242b (Indicative Value 0.35 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.35 | 4
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 242b (Indicative Value 0.46 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.46 | 32
| Mean |  | 0.455 | 16
| Median |  | 0.455 | 0.455
| Std Dev. |  | 0.0070711 | 38
| Rel.Std.Dev. |  | 1.554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 242b (Indicative Value 276 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 275 | 1
| 1 | 2 | 277 | 33
| Mean |  | 276 | 16
| Median |  | 276 | 276
| Std Dev. |  | 1.4142 | 39
| Rel.Std.Dev. |  | 0.5124%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 242b (Indicative Value 16.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.5 | 1
| 1 | 2 | 16.5 | 34
| Mean |  | 16.5 | 16
| Median |  | 16.5 | 16.5
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 242b (Indicative Value 20.7 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.7 | 1
| 1 | 2 | 20.7 | 35
| Mean |  | 20.7 | 16
| Median |  | 20.7 | 20.7
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 242b (Indicative Value 2.32 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.32 | 1
| 1 | 2 | 2.31 | 5
| Mean |  | 2.315 | 16
| Median |  | 2.315 | 2.315
| Std Dev. |  | 0.0070711 | 42
| Rel.Std.Dev. |  | 0.3054%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 242b (Indicative Value 118 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 115 | 1
| 1 | 2 | 120 | 15
| Mean |  | 117.5 | 16
| Median |  | 117.5 | 117.5
| Std Dev. |  | 3.5355 | 43
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 242b (Indicative Value 70 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 69.5 | 16
| Mean |  | 69.75 | 16
| Median |  | 69.75 | 69.75
| Std Dev. |  | 0.35355 | 44
| Rel.Std.Dev. |  | 0.5069%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.