## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 262b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 104.108 | 101.2475 | 106.9685 | 101.8939 | 106.3221
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppb) | 93.5224 | 86.762 | 100.2829 | 91.3436 | 95.7012
Cyanide Leach
Au, Gold (ppb) | 82.1072 | 74.644 | 89.5705 | 81.4309 | 82.7835
PhotonAssay (recommended gross mass* 460±20 g)
Au, Gold (ppb) | 108.0455 | 101.9701 | 114.1208 | 107.3727 | 108.7182
4-Acid Digestion
Ag, Silver (ppm) | 0.15767 | 0.12527 | 0.19007 | IND | IND
Al, Aluminium (wt.%) | 5.0095 | 4.8483 | 5.1707 | 4.9097 | 5.1093
As, Arsenic (ppm) | 30.2542 | 27.9813 | 32.5271 | 28.5911 | 31.9173
Ba, Barium (ppm) | 336.4853 | 325.4717 | 347.499 | 330.1623 | 342.8084
Be, Beryllium (ppm) | 1.5178 | 1.4617 | 1.5739 | 1.475 | 1.5606
Bi, Bismuth (ppm) | 0.73359 | 0.6815 | 0.78569 | 0.7064 | 0.76079
Ca, Calcium (wt.%) | 0.38385 | 0.3695 | 0.3982 | 0.37438 | 0.39331
Cd, Cadmium (ppm) | 0.061625 | 0.043418 | 0.079831 | 0.054929 | 0.06832
Ce, Cerium (ppm) | 67.7365 | 64.7956 | 70.6774 | 65.7438 | 69.7292
Co, Cobalt (ppm) | 5.7369 | 5.4147 | 6.0592 | 5.562 | 5.9118
Cr, Chromium (ppm) | 49.6568 | 46.3452 | 52.9685 | 47.6115 | 51.7021
Cs, Caesium (ppm) | 5.0194 | 4.7929 | 5.2459 | 4.8794 | 5.1594
Cu, Copper (ppm) | 70.2385 | 67.1717 | 73.3054 | 67.5592 | 72.9178
Dy, Dysprosium (ppm) | 2.3735 | 1.944 | 2.8029 | 2.2149 | 2.532
Er, Erbium (ppm) | 1.307 | 1.1431 | 1.4709 | 1.211 | 1.403
Eu, Europium (ppm) | 0.73702 | 0.6794 | 0.79464 | 0.69231 | 0.78172
Fe, Iron (wt.%) | 2.4603 | 2.3894 | 2.5311 | 2.4137 | 2.5069
Ga, Gallium (ppm) | 12.5909 | 11.9227 | 13.2591 | 12.2274 | 12.9545
Gd, Gadolinium (ppm) | 3.3522 | 2.9159 | 3.7885 | 3.1407 | 3.5637
Hf, Hafnium (ppm) | 2.5356 | 2.3223 | 2.7489 | 2.3943 | 2.6769
Ho, Holmium (ppm) | 0.43225 | 0.37202 | 0.49248 | 0.39295 | 0.47154
In, Indium (ppm) | 0.063365 | 0.055756 | 0.070973 | 0.057309 | 0.069421
K, Potassium (wt.%) | 1.6341 | 1.5717 | 1.6965 | 1.5937 | 1.6745
La, Lanthanum (ppm) | 34.5743 | 32.8192 | 36.3294 | 33.3572 | 35.7914
Li, Lithium (ppm) | 34.8551 | 33.6386 | 36.0717 | 33.857 | 35.8533
Lu, Lutetium (ppm) | 0.19321 | 0.15431 | 0.23211 | 0.17789 | 0.20853
Mg, Magnesium (wt.%) | 0.55673 | 0.53747 | 0.57599 | 0.54429 | 0.56916
Mn, Manganese (wt.%) | 0.01868 | 0.018018 | 0.019341 | 0.018268 | 0.019091
Mo, Molybdenum (ppm) | 1.4827 | 1.2824 | 1.6831 | 1.3364 | 1.6291
Na, Sodium (wt.%) | 0.18978 | 0.18145 | 0.19811 | 0.18363 | 0.19593
Nb, Niobium (ppm) | 7.704 | 6.8631 | 8.545 | 7.2088 | 8.1993
Nd, Neodymium (ppm) | 23.9658 | 18.8827 | 29.0488 | 23.274 | 24.6575
Ni, Nickel (ppm) | 22.0385 | 20.3517 | 23.7253 | 21.1352 | 22.9419
P, Phosphorus (wt.%) | 0.023557 | 0.022332 | 0.024782 | 0.022836 | 0.024278
Pb, Lead (ppm) | 19.354 | 18.5307 | 20.1774 | 18.6754 | 20.0327
Pr, Praseodymium (ppm) | 6.9687 | 6.13 | 7.8074 | 6.6534 | 7.284
Rb, Rubidium (ppm) | 110.1437 | 104.7819 | 115.5056 | 106.6235 | 113.664
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.036377 | 0.032862 | 0.039891 | 0.034582 | 0.038172
Sb, Antimony (ppm) | 5.4055 | 5.0852 | 5.7257 | 5.1336 | 5.6774
Sc, Scandium (ppm) | 7.0173 | 6.5941 | 7.4405 | 6.776 | 7.2586
Sm, Samarium (ppm) | 4.3519 | 3.8805 | 4.8234 | 4.1275 | 4.5763
Sn, Tin (ppm) | 6.7236 | 6.3961 | 7.051 | 6.4055 | 7.0417
Sr, Strontium (ppm) | 44.0777 | 41.8725 | 46.2829 | 43.1521 | 45.0033
Ta, Tantalum (ppm) | 0.78102 | 0.679 | 0.88303 | 0.73309 | 0.82894
Tb, Terbium (ppm) | 0.44886 | 0.392 | 0.50573 | 0.4177 | 0.48003
Te, Tellurium (ppm) | 0.11035 | 0.077758 | 0.14295 | IND | IND
Th, Thorium (ppm) | 12.6147 | 11.9018 | 13.3275 | 12.3442 | 12.8851
Ti, Titanium (wt.%) | 0.2525 | 0.23774 | 0.26726 | 0.24297 | 0.26203
Tl, Thallium (ppm) | 0.88233 | 0.83347 | 0.9312 | 0.8399 | 0.92476
Tm, Thulium (ppm) | 0.17789 | 0.1569 | 0.19888 | IND | IND
U, Uranium (ppm) | 1.9468 | 1.8264 | 2.0672 | 1.8812 | 2.0124
V, Vanadium (ppm) | 51.3458 | 48.9758 | 53.7158 | 49.6766 | 53.0151
W, Tungsten (ppm) | 3.0698 | 2.7962 | 3.3433 | 2.8423 | 3.2972
Y, Yttrium (ppm) | 12.0098 | 11.1384 | 12.8812 | 11.5616 | 12.458
Yb, Ytterbium (ppm) | 1.2977 | 1.1147 | 1.4806 | 1.2479 | 1.3475
Zn, Zinc (ppm) | 47.1602 | 45.0943 | 49.2261 | 46.0593 | 48.2611
Zr, Zirconium (ppm) | 89.1404 | 83.3703 | 94.9104 | 86.7524 | 91.5283
Aqua Regia Digestion
Ag, Silver (ppm) | 0.13212 | 0.1177 | 0.14655 | IND | IND
Al, Aluminium (wt.%) | 0.92235 | 0.82083 | 1.0239 | 0.88812 | 0.95657
As, Arsenic (ppm) | 13.8769 | 12.9663 | 14.7874 | 13.4118 | 14.342
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 74.5221 | 69.3542 | 79.69 | 71.8002 | 77.244
Be, Beryllium (ppm) | 0.47075 | 0.42999 | 0.51151 | 0.46002 | 0.48149
Bi, Bismuth (ppm) | 0.68632 | 0.64161 | 0.73102 | 0.65323 | 0.71941
Ca, Calcium (wt.%) | 0.26614 | 0.25616 | 0.27612 | 0.25576 | 0.27652
Cd, Cadmium (ppm) | 0.062926 | 0.048864 | 0.076988 | IND | IND
Ce, Cerium (ppm) | 31.2757 | 29.3867 | 33.1646 | 30.2447 | 32.3066
Co, Cobalt (ppm) | 4.0244 | 3.7082 | 4.3405 | 3.8949 | 4.1538
Cr, Chromium (ppm) | 31.8 | 29.9373 | 33.6627 | 30.8256 | 32.7745
Cs, Caesium (ppm) | 1.6534 | 1.4729 | 1.8339 | 1.5995 | 1.7073
Cu, Copper (ppm) | 66.1824 | 63.741 | 68.6238 | 64.7147 | 67.65
Fe, Iron (wt.%) | 2.0674 | 1.9914 | 2.1434 | 2.0233 | 2.1115
Ga, Gallium (ppm) | 2.9263 | 2.5271 | 3.3256 | 2.7762 | 3.0764
Hf, Hafnium (ppm) | 0.32979 | 0.29026 | 0.36933 | 0.31074 | 0.34884
Hg, Mercury (ppm) | 0.055583 | 0.04254 | 0.068626 | IND | IND
In, Indium (ppm) | 0.019342 | 0.016129 | 0.022555 | IND | IND
K, Potassium (wt.%) | 0.21649 | 0.1987 | 0.23428 | 0.2095 | 0.22348
La, Lanthanum (ppm) | 14.489 | 13.6907 | 15.2873 | 14.0944 | 14.8836
Li, Lithium (ppm) | 11.3032 | 10.1717 | 12.4346 | 10.9787 | 11.6277
Mg, Magnesium (wt.%) | 0.28656 | 0.25993 | 0.31319 | 0.27687 | 0.29625
Mn, Manganese (wt.%) | 0.014201 | 0.013576 | 0.014827 | 0.013838 | 0.014564
Mo, Molybdenum (ppm) | 1.3885 | 1.2429 | 1.534 | 1.2732 | 1.5038
Ni, Nickel (ppm) | 15.9961 | 14.6728 | 17.3194 | 15.405 | 16.5873
P, Phosphorus (wt.%) | 0.019492 | 0.01849 | 0.020493 | 0.018836 | 0.020148
Pb, Lead (ppm) | 16.6464 | 15.7836 | 17.5093 | 16.0439 | 17.2489
Rb, Rubidium (ppm) | 16.9634 | 15.8938 | 18.033 | 16.4154 | 17.5113
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.030629 | 0.029652 | 0.031606 | 0.029265 | 0.031993
Sb, Antimony (ppm) | 3.5503 | 3.2386 | 3.862 | 3.3429 | 3.7577
Sc, Scandium (ppm) | 1.5522 | 1.4028 | 1.7016 | IND | IND
Sn, Tin (ppm) | 1.6649 | 1.4963 | 1.8334 | 1.577 | 1.7528
Sr, Strontium (ppm) | 12.0272 | 11.2876 | 12.7668 | 11.5582 | 12.4963
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.25911 | 0.22264 | 0.29557 | IND | IND
Te, Tellurium (ppm) | 0.089796 | 0.068181 | 0.11141 | IND | IND
Th, Thorium (ppm) | 8.8889 | 8.3409 | 9.4369 | 8.5528 | 9.225
Ti, Titanium (wt.%) | 0.017409 | 0.015492 | 0.019325 | 0.016396 | 0.018421
Tl, Thallium (ppm) | 0.27935 | 0.25722 | 0.30147 | 0.26322 | 0.29547
U, Uranium (ppm) | 1.1255 | 1.0622 | 1.1888 | 1.0929 | 1.1581
V, Vanadium (ppm) | 19.2988 | 18.3341 | 20.2634 | 18.441 | 20.1565
W, Tungsten (ppm) | 0.64274 | 0.52715 | 0.75834 | 0.60344 | 0.68205
Y, Yttrium (ppm) | 5.0566 | 4.7473 | 5.366 | 4.9424 | 5.1709
Yb, Ytterbium (ppm) | 0.43635 | 0.33799 | 0.5347 | IND | IND
Zn, Zinc (ppm) | 28.4996 | 26.3637 | 30.6356 | 27.6731 | 29.3262
Zr, Zirconium (ppm) | 13.3798 | 11.9091 | 14.8506 | 12.8968 | 13.8629
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 262b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 5 | Pt | ppb | < 10
4-Acid Digestion
Ge | ppm | 0.12766 | Hg | ppm | 0.084283 | Se | ppm | 0.72622
Aqua Regia Digestion
Dy | ppm | 1.1762 | Lu | ppm | 0.059761 | Pt | ppb | < 5
Er | ppm | 0.49935 | Na | wt.% | 0.01001 | Se | ppm | 0.2475
Eu | ppm | 0.43138 | Nb | ppm | 0.16931 | Sm | ppm | 2.6196
Gd | ppm | 2.0099 | Nd | ppm | 14.232 | Tm | ppm | 0.067115
Ge | ppm | 0.059708 | Pd | ppb | < 10
Ho | ppm | 0.19434 | Pr | ppm | 3.8074
Borate Fusion XRF
Al2O3 | wt.% | 9.58 | MgO | wt.% | 0.945 | SiO2 | wt.% | 80.025
CaO | wt.% | 0.49 | MnO | wt.% | 0.023 | SO3 | wt.% | 0.0915
Fe2O3 | wt.% | 3.51 | Na2O | wt.% | 0.26 | TiO2 | wt.% | 0.4895
K2O | wt.% | 1.97 | P2O5 | wt.% | 0.057
Thermogravimetry
LOI1000 | wt.% | 2.445
Infrared Combustion
C | wt.% | 0.145 | S | wt.% | 0.02
Laser Ablation ICP-MS
Ag | ppm | 0.15 | Hf | ppm | 6.74 | Sm | ppm | 4.68
As | ppm | 29.7 | Ho | ppm | 0.775 | Sn | ppm | 9.9
Ba | ppm | 325 | In | ppm | 0.05 | Sr | ppm | 41.9
Be | ppm | 1.9 | La | ppm | 34.2 | Ta | ppm | 0.93
Bi | ppm | 0.79 | Lu | ppm | 0.335 | Tb | ppm | 0.625
Cd | ppm | 0.125 | Mn | wt.% | 0.0184 | Te | ppm | < 0.2
Ce | ppm | 65.95 | Mo | ppm | 1.7 | Th | ppm | 12.9
Co | ppm | 6.1 | Nb | ppm | 10.4 | Ti | wt.% | 0.29
Cr | ppm | 57 | Nd | ppm | 26.6 | Tl | ppm | 0.8
Cs | ppm | 4.85 | Ni | ppm | 22 | Tm | ppm | 0.325
Cu | ppm | 71 | Pb | ppm | 20 | U | ppm | 2.185
Dy | ppm | 3.825 | Pr | ppm | 7.655 | V | ppm | 54
Er | ppm | 2.235 | Rb | ppm | 104 | W | ppm | 3.25
Eu | ppm | 0.785 | Re | ppm | < 0.01 | Y | ppm | 20.65
Ga | ppm | 11.8 | Sb | ppm | 6.25 | Yb | ppm | 2.225
Gd | ppm | 3.885 | Sc | ppm | 6.8 | Zn | ppm | 47.5
Ge | ppm | 1.25 | Se | ppm | < 5 | Zr | ppm | 233.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 262b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 104.108 | 6.0372 | 92.0336 | 116.1824 | 85.9964 | 122.2196 | 5.799% | 11.6% | 17.4% | 98.9026 | 109.3134
Aqua Regia Digestion (sample weights 10-50g)
Au, ppb | 93.5224 | 13.5549 | 66.4126 | 120.6323 | 52.8577 | 134.1872 | 14.49% | 28.99% | 43.48% | 88.8463 | 98.1986
Cyanide Leach
Au, ppb | 82.1072 | 14.5316 | 53.0441 | 111.1703 | 38.5126 | 125.7019 | 17.7% | 35.4% | 53.09% | 78.0019 | 86.2126
PhotonAssay
Au, ppb | 108.0455 | 23.3689 | 61.3076 | 154.7833 | 37.9387 | 178.1522 | 21.63% | 43.26% | 64.89% | 102.6432 | 113.4477
4-Acid Digestion
Ag, ppm | 0.15767 | 0.019416 | 0.11883 | 0.1965 | 0.099419 | 0.21591 | 12.31% | 24.63% | 36.94% | 0.14978 | 0.16555
Al, wt.% | 5.0095 | 0.21417 | 4.5811 | 5.4378 | 4.3669 | 5.652 | 4.275% | 8.551% | 12.83% | 4.759 | 5.2599
As, ppm | 30.2542 | 2.9504 | 24.3535 | 36.1549 | 21.4032 | 39.1053 | 9.752% | 19.5% | 29.26% | 28.7415 | 31.7669
Ba, ppm | 336.4853 | 11.3069 | 313.8716 | 359.0991 | 302.5647 | 370.406 | 3.36% | 6.721% | 10.08% | 319.6611 | 353.3096
Be, ppm | 1.5178 | 0.033414 | 1.451 | 1.5846 | 1.4176 | 1.618 | 2.201% | 4.403% | 6.604% | 1.4419 | 1.5937
Bi, ppm | 0.73359 | 0.050858 | 0.63188 | 0.83531 | 0.58102 | 0.88617 | 6.933% | 13.87% | 20.8% | 0.69691 | 0.77027
Ca, wt.% | 0.38385 | 0.013953 | 0.35594 | 0.41175 | 0.34199 | 0.42571 | 3.635% | 7.27% | 10.91% | 0.36465 | 0.40304
Cd, ppm | 0.061625 | 0.016877 | 0.027871 | 0.095378 | 0.010994 | 0.11226 | 27.39% | 54.77% | 82.16% | 0.058544 | 0.064706
Ce, ppm | 67.7365 | 2.5587 | 62.6191 | 72.8539 | 60.0604 | 75.4126 | 3.777% | 7.555% | 11.33% | 64.3497 | 71.1233
Co, ppm | 5.7369 | 0.3493 | 5.0383 | 6.4355 | 4.689 | 6.7848 | 6.089% | 12.18% | 18.27% | 5.4501 | 6.0238
Cr, ppm | 49.6568 | 4.9913 | 39.6742 | 59.6395 | 34.6829 | 64.6308 | 10.05% | 20.1% | 30.15% | 47.174 | 52.1397
Cs, ppm | 5.0194 | 0.16925 | 4.6809 | 5.3579 | 4.5117 | 5.5271 | 3.372% | 6.744% | 10.12% | 4.7684 | 5.2704
Cu, ppm | 70.2385 | 3.465 | 63.3085 | 77.1685 | 59.8435 | 80.6335 | 4.933% | 9.866% | 14.8% | 66.7266 | 73.7504
Dy, ppm | 2.3735 | 0.38854 | 1.5964 | 3.1505 | 1.2078 | 3.5391 | 16.37% | 32.74% | 49.11% | 2.2548 | 2.4921
Er, ppm | 1.307 | 0.12831 | 1.0504 | 1.5636 | 0.92208 | 1.6919 | 9.817% | 19.63% | 29.45% | 1.2417 | 1.3724
Eu, ppm | 0.73702 | 0.028031 | 0.68096 | 0.79308 | 0.65292 | 0.82111 | 3.803% | 7.607% | 11.41% | 0.70017 | 0.77387
Fe, wt.% | 2.4603 | 0.049127 | 2.362 | 2.5585 | 2.3129 | 2.6077 | 1.997% | 3.994% | 5.99% | 2.3373 | 2.5833
Ga, ppm | 12.5909 | 0.71942 | 11.1521 | 14.0298 | 10.4327 | 14.7492 | 5.714% | 11.43% | 17.14% | 11.9614 | 13.2205
Gd, ppm | 3.3522 | 0.25545 | 2.8413 | 3.8631 | 2.5859 | 4.1186 | 7.62% | 15.24% | 22.86% | 3.1846 | 3.5198
Hf, ppm | 2.5356 | 0.23793 | 2.0597 | 3.0114 | 1.8218 | 3.2494 | 9.383% | 18.77% | 28.15% | 2.4088 | 2.6624
Ho, ppm | 0.43225 | 0.047656 | 0.33694 | 0.52756 | 0.28928 | 0.57522 | 11.03% | 22.05% | 33.08% | 0.41064 | 0.45386
In, ppm | 0.063365 | 0.0058687 | 0.051627 | 0.075102 | 0.045759 | 0.080971 | 9.262% | 18.52% | 27.79% | 0.060196 | 0.066533
K, wt.% | 1.6341 | 0.082021 | 1.4701 | 1.7982 | 1.3881 | 1.8802 | 5.019% | 10.04% | 15.06% | 1.5524 | 1.7158
La, ppm | 34.5743 | 1.2179 | 32.1386 | 37.01 | 30.9208 | 38.2279 | 3.522% | 7.045% | 10.57% | 32.8456 | 36.303
Li, ppm | 34.8551 | 1.2703 | 32.3145 | 37.3958 | 31.0442 | 38.6661 | 3.645% | 7.289% | 10.93% | 33.1124 | 36.5979
Lu, ppm | 0.19321 | 0.024864 | 0.14348 | 0.24294 | 0.11862 | 0.2678 | 12.87% | 25.74% | 38.61% | 0.18355 | 0.20287
Mg, wt.% | 0.55673 | 0.014362 | 0.528 | 0.58545 | 0.51364 | 0.59981 | 2.58% | 5.159% | 7.739% | 0.52889 | 0.58456
Mn, wt.% | 0.01868 | 0.00076788 | 0.017144 | 0.020215 | 0.016376 | 0.020983 | 4.111% | 8.222% | 12.33% | 0.017746 | 0.019614
Mo, ppm | 1.4827 | 0.12068 | 1.2414 | 1.7241 | 1.1207 | 1.8448 | 8.139% | 16.28% | 24.42% | 1.4086 | 1.5569
Na, wt.% | 0.18978 | 0.011053 | 0.16767 | 0.21188 | 0.15662 | 0.22294 | 5.824% | 11.65% | 17.47% | 0.18029 | 0.19927
Nb, ppm | 7.704 | 1.301 | 5.1021 | 10.306 | 3.8011 | 11.607 | 16.89% | 33.77% | 50.66% | 7.3188 | 8.0892
Nd, ppm | 23.9658 | 4.3903 | 15.1851 | 32.7465 | 10.7947 | 37.1368 | 18.32% | 36.64% | 54.96% | 22.7675 | 25.1641
Ni, ppm | 22.0385 | 1.6608 | 18.7168 | 25.3602 | 17.056 | 27.021 | 7.536% | 15.07% | 22.61% | 20.9366 | 23.1404
P, wt.% | 0.023557 | 0.0011681 | 0.021221 | 0.025893 | 0.020053 | 0.027061 | 4.959% | 9.917% | 14.88% | 0.022379 | 0.024735
Pb, ppm | 19.354 | 0.72605 | 17.9019 | 20.8061 | 17.1759 | 21.5322 | 3.751% | 7.503% | 11.25% | 18.3863 | 20.3217
Pr, ppm | 6.9687 | 0.71869 | 5.5313 | 8.4061 | 4.8126 | 9.1247 | 10.31% | 20.63% | 30.94% | 6.6202 | 7.3171
Rb, ppm | 110.1437 | 6.2324 | 97.6788 | 122.6086 | 91.4464 | 128.841 | 5.658% | 11.32% | 16.98% | 104.6365 | 115.6509
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.036377 | 0.005554 | 0.025269 | 0.047485 | 0.019715 | 0.053039 | 15.27% | 30.54% | 45.8% | 0.034558 | 0.038196
Sb, ppm | 5.4055 | 0.33409 | 4.7373 | 6.0737 | 4.4032 | 6.4077 | 6.181% | 12.36% | 18.54% | 5.1352 | 5.6757
Sc, ppm | 7.0173 | 0.44719 | 6.1229 | 7.9117 | 5.6757 | 8.3589 | 6.373% | 12.75% | 19.12% | 6.6664 | 7.3681
Sm, ppm | 4.3519 | 0.34814 | 3.6556 | 5.0482 | 3.3075 | 5.3964 | 8% | 16% | 24% | 4.1343 | 4.5695
Sn, ppm | 6.7236 | 0.29191 | 6.1398 | 7.3074 | 5.8479 | 7.5993 | 4.342% | 8.683% | 13.02% | 6.3874 | 7.0598
Sr, ppm | 44.0777 | 1.6733 | 40.731 | 47.4244 | 39.0577 | 49.0977 | 3.796% | 7.593% | 11.39% | 41.8738 | 46.2816
Ta, ppm | 0.78102 | 0.12096 | 0.5391 | 1.0229 | 0.41814 | 1.1439 | 15.49% | 30.97% | 46.46% | 0.74197 | 0.82007
Tb, ppm | 0.44886 | 0.037291 | 0.37428 | 0.52345 | 0.33699 | 0.56074 | 8.308% | 16.62% | 24.92% | 0.42642 | 0.47131
Te, ppm | 0.11035 | 0.020347 | 0.069658 | 0.15105 | 0.049311 | 0.17139 | 18.44% | 36.88% | 55.31% | 0.10483 | 0.11587
Th, ppm | 12.6147 | 0.64246 | 11.3297 | 13.8996 | 10.6873 | 14.542 | 5.093% | 10.19% | 15.28% | 11.9839 | 13.2454
Ti, wt.% | 0.2525 | 0.019299 | 0.2139 | 0.2911 | 0.1946 | 0.3104 | 7.643% | 15.29% | 22.93% | 0.23988 | 0.26513
Tl, ppm | 0.88233 | 0.047018 | 0.7883 | 0.97637 | 0.74128 | 1.0234 | 5.329% | 10.66% | 15.99% | 0.83822 | 0.92645
Tm, ppm | 0.17789 | 0.017168 | 0.14355 | 0.21222 | 0.12638 | 0.22939 | 9.651% | 19.3% | 28.95% | 0.16899 | 0.18678
U, ppm | 1.9468 | 0.097306 | 1.7522 | 2.1414 | 1.6549 | 2.2387 | 4.998% | 9.997% | 14.99% | 1.8495 | 2.0441
V, ppm | 51.3458 | 2.679 | 45.9879 | 56.7038 | 43.3089 | 59.3828 | 5.218% | 10.44% | 15.65% | 48.7785 | 53.9131
W, ppm | 3.0698 | 0.28388 | 2.502 | 3.6375 | 2.2181 | 3.9214 | 9.248% | 18.5% | 27.74% | 2.9163 | 3.2232
Y, ppm | 12.0098 | 0.94999 | 10.1098 | 13.9098 | 9.1598 | 14.8598 | 7.91% | 15.82% | 23.73% | 11.4093 | 12.6103
Yb, ppm | 1.2977 | 0.18753 | 0.9226 | 1.6727 | 0.73507 | 1.8603 | 14.45% | 28.9% | 43.35% | 1.2328 | 1.3625
Zn, ppm | 47.1602 | 2.2824 | 42.5953 | 51.725 | 40.3129 | 54.0075 | 4.84% | 9.68% | 14.52% | 44.8022 | 49.5182
Zr, ppm | 89.1404 | 6.9585 | 75.2234 | 103.0573 | 68.265 | 110.0157 | 7.806% | 15.61% | 23.42% | 84.6833 | 93.5974
Aqua Regia Digestion
Ag, ppm | 0.13212 | 0.0080203 | 0.11608 | 0.14817 | 0.10806 | 0.15619 | 6.07% | 12.14% | 18.21% | 0.12552 | 0.13873
Al, wt.% | 0.92235 | 0.16975 | 0.58285 | 1.2618 | 0.4131 | 1.4316 | 18.4% | 36.81% | 55.21% | 0.87623 | 0.96846
As, ppm | 13.8769 | 0.76005 | 12.3568 | 15.397 | 11.5967 | 16.157 | 5.477% | 10.95% | 16.43% | 13.183 | 14.5707
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 74.5221 | 4.1375 | 66.2471 | 82.7971 | 62.1096 | 86.9346 | 5.552% | 11.1% | 16.66% | 70.796 | 78.2482
Be, ppm | 0.47075 | 0.049475 | 0.3718 | 0.5697 | 0.32233 | 0.61918 | 10.51% | 21.02% | 31.53% | 0.44722 | 0.49429
Bi, ppm | 0.68632 | 0.042807 | 0.6007 | 0.77193 | 0.5579 | 0.81474 | 6.237% | 12.47% | 18.71% | 0.652 | 0.72063
Ca, wt.% | 0.26614 | 0.010743 | 0.24465 | 0.28763 | 0.23391 | 0.29837 | 4.036% | 8.073% | 12.11% | 0.25283 | 0.27945
Cd, ppm | 0.062926 | 0.010072 | 0.042781 | 0.083071 | 0.032709 | 0.093143 | 16.01% | 32.01% | 48.02% | 0.05978 | 0.066072
Ce, ppm | 31.2757 | 2.0416 | 27.1925 | 35.3588 | 25.151 | 37.4003 | 6.528% | 13.06% | 19.58% | 29.7119 | 32.8394
Co, ppm | 4.0244 | 0.43223 | 3.1599 | 4.8888 | 2.7277 | 5.321 | 10.74% | 21.48% | 32.22% | 3.8231 | 4.2256
Cr, ppm | 31.8 | 1.4321 | 28.9358 | 34.6643 | 27.5037 | 36.0964 | 4.504% | 9.007% | 13.51% | 30.21 | 33.39
Cs, ppm | 1.6534 | 0.27331 | 1.1068 | 2.2 | 0.83345 | 2.4733 | 16.53% | 33.06% | 49.59% | 1.5707 | 1.7361
Cu, ppm | 66.1824 | 2.4209 | 61.3405 | 71.0242 | 58.9196 | 73.4452 | 3.658% | 7.316% | 10.97% | 62.8733 | 69.4915
Fe, wt.% | 2.0674 | 0.12208 | 1.8232 | 2.3115 | 1.7011 | 2.4336 | 5.905% | 11.81% | 17.72% | 1.964 | 2.1707
Ga, ppm | 2.9263 | 0.53135 | 1.8636 | 3.989 | 1.3323 | 4.5204 | 18.16% | 36.31% | 54.47% | 2.78 | 3.0726
Hf, ppm | 0.32979 | 0.030707 | 0.26838 | 0.39121 | 0.23767 | 0.42191 | 9.311% | 18.62% | 27.93% | 0.3133 | 0.34628
Hg, ppm | 0.055583 | 0.012973 | 0.029638 | 0.081529 | 0.016666 | 0.094501 | 23.34% | 46.68% | 70.02% | 0.052804 | 0.058362
In, ppm | 0.019342 | 0.0019143 | 0.015513 | 0.02317 | 0.013599 | 0.025085 | 9.897% | 19.79% | 29.69% | 0.018375 | 0.020309
K, wt.% | 0.21649 | 0.027764 | 0.16096 | 0.27202 | 0.1332 | 0.29978 | 12.82% | 25.65% | 38.47% | 0.20567 | 0.22732
La, ppm | 14.489 | 1.13 | 12.229 | 16.749 | 11.099 | 17.879 | 7.799% | 15.6% | 23.4% | 13.7645 | 15.2134
Li, ppm | 11.3032 | 1.6829 | 7.9374 | 14.6689 | 6.2546 | 16.3517 | 14.89% | 29.78% | 44.67% | 10.738 | 11.8683
Mg, wt.% | 0.28656 | 0.046919 | 0.19272 | 0.3804 | 0.1458 | 0.42732 | 16.37% | 32.75% | 49.12% | 0.27223 | 0.30089
Mn, wt.% | 0.014201 | 0.0007065 | 0.012788 | 0.015614 | 0.012082 | 0.016321 | 4.975% | 9.95% | 14.92% | 0.013491 | 0.014911
Mo, ppm | 1.3885 | 0.14877 | 1.0909 | 1.686 | 0.94217 | 1.8348 | 10.71% | 21.43% | 32.14% | 1.319 | 1.4579
Ni, ppm | 15.9961 | 1.1463 | 13.7034 | 18.2888 | 12.5571 | 19.4351 | 7.166% | 14.33% | 21.5% | 15.1963 | 16.7959
P, wt.% | 0.019492 | 0.00085206 | 0.017788 | 0.021196 | 0.016936 | 0.022048 | 4.371% | 8.743% | 13.11% | 0.018517 | 0.020466
Pb, ppm | 16.6464 | 0.9525 | 14.7414 | 18.5514 | 13.7889 | 19.5039 | 5.722% | 11.44% | 17.17% | 15.8141 | 17.4788
Rb, ppm | 16.9634 | 1.0718 | 14.8198 | 19.107 | 13.7479 | 20.1788 | 6.318% | 12.64% | 18.96% | 16.1152 | 17.8116
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.030629 | 0.0012087 | 0.028212 | 0.033046 | 0.027003 | 0.034255 | 3.946% | 7.892% | 11.84% | 0.029098 | 0.032161
Sb, ppm | 3.5503 | 0.48246 | 2.5854 | 4.5152 | 2.1029 | 4.9977 | 13.59% | 27.18% | 40.77% | 3.3728 | 3.7278
Sc, ppm | 1.5522 | 0.11281 | 1.3266 | 1.7778 | 1.2138 | 1.8907 | 7.268% | 14.54% | 21.8% | 1.4746 | 1.6298
Sn, ppm | 1.6649 | 0.21563 | 1.2336 | 2.0961 | 1.018 | 2.3118 | 12.95% | 25.9% | 38.86% | 1.5816 | 1.7481
Sr, ppm | 12.0272 | 0.72346 | 10.5803 | 13.4741 | 9.8568 | 14.1976 | 6.015% | 12.03% | 18.05% | 11.4259 | 12.6286
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.25911 | 0.032993 | 0.19312 | 0.32509 | 0.16013 | 0.35809 | 12.73% | 25.47% | 38.2% | 0.24615 | 0.27206
Te, ppm | 0.089796 | 0.017462 | 0.054873 | 0.12472 | 0.037411 | 0.14218 | 19.45% | 38.89% | 58.34% | 0.085307 | 0.094286
Th, ppm | 8.8889 | 0.58307 | 7.7227 | 10.055 | 7.1397 | 10.6381 | 6.56% | 13.12% | 19.68% | 8.4444 | 9.3333
Ti, wt.% | 0.017409 | 0.0028277 | 0.011753 | 0.023064 | 0.0089254 | 0.025892 | 16.24% | 32.49% | 48.73% | 0.016538 | 0.018279
Tl, ppm | 0.27935 | 0.022471 | 0.23441 | 0.32429 | 0.21194 | 0.34676 | 8.044% | 16.09% | 24.13% | 0.26538 | 0.29332
U, ppm | 1.1255 | 0.086969 | 0.95156 | 1.2994 | 0.86459 | 1.3864 | 7.727% | 15.45% | 23.18% | 1.0692 | 1.1818
V, ppm | 19.2988 | 1.036 | 17.2267 | 21.3708 | 16.1907 | 22.4068 | 5.368% | 10.74% | 16.1% | 18.3338 | 20.2637
W, ppm | 0.64274 | 0.12215 | 0.39845 | 0.88703 | 0.27631 | 1.0092 | 19% | 38.01% | 57.01% | 0.61061 | 0.67488
Y, ppm | 5.0566 | 0.41561 | 4.2254 | 5.8879 | 3.8098 | 6.3035 | 8.219% | 16.44% | 24.66% | 4.8038 | 5.3095
Yb, ppm | 0.43635 | 0.04641 | 0.34353 | 0.52917 | 0.29712 | 0.57558 | 10.64% | 21.27% | 31.91% | 0.41453 | 0.45817
Zn, ppm | 28.4996 | 3.4535 | 21.5926 | 35.4067 | 18.139 | 38.8602 | 12.12% | 24.24% | 36.35% | 27.0746 | 29.9246
Zr, ppm | 13.3798 | 2.0052 | 9.3695 | 17.3902 | 7.3643 | 19.3954 | 14.99% | 29.97% | 44.96% | 12.7109 | 14.0488
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 262b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 262b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 262b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppb) | Upscaled Value (ppb)
2 | 1 | ANSLu | 1 | 14 | 7 | 245380 | 0.082949 | 99.4465 | 107.0578
2 | 1 | ANSLu | 1 | 5 | 3 | 245381 | 0.087825 | 95.5319 | 106.8485
2 | 1 | ANSLu | 1 | 2 | 1 | 245382 | 0.08609 | 100.5888 | 107.1189
2 | 1 | ANSLu | 1 | 9 | 5 | 245383 | 0.083469 | 124.6022 | 108.4031
2 | 1 | ANSLu | 1 | 6 | 3 | 245384 | 0.084826 | 100.9293 | 107.1371
2 | 1 | ANSLu | 1 | 13 | 7 | 245385 | 0.084852 | 141.4018 | 109.3016
2 | 1 | ANSLu | 1 | 7 | 4 | 245386 | 0.085069 | 113.0539 | 107.7855
2 | 1 | ANSLu | 1 | 20 | 10 | 245387 | 0.087489 | 92.3597 | 106.6788
2 | 1 | ANSLu | 1 | 19 | 10 | 245388 | 0.086695 | 134.3882 | 108.9265
2 | 1 | ANSLu | 1 | 11 | 6 | 245389 | 0.083885 | 102.734 | 107.2336
2 | 1 | ANSLu | 1 | 10 | 5 | 245390 | 0.087055 | 101.4569 | 107.1653
2 | 1 | ANSLu | 1 | 8 | 4 | 245391 | 0.088159 | 94.5215 | 106.7944
2 | 1 | ANSLu | 1 | 16 | 8 | 245392 | 0.086328 | 97.8839 | 106.9743
2 | 1 | ANSLu | 1 | 18 | 9 | 245393 | 0.086263 | 115.088 | 107.8943
2 | 1 | ANSLu | 1 | 17 | 9 | 245394 | 0.083795 | 109.7681 | 107.6098
2 | 1 | ANSLu | 1 | 15 | 8 | 245395 | 0.086884 | 120.185 | 108.1669
2 | 1 | ANSLu | 1 | 4 | 2 | 245396 | 0.085991 | 106.0404 | 107.4105
2 | 1 | ANSLu | 1 | 3 | 2 | 245397 | 0.085301 | 102.8113 | 107.2378
2 | 1 | ANSLu | 1 | 12 | 6 | 245398 | 0.086537 | 98.0939 | 106.9855
2 | 1 | ANSLu | 1 | 1 | 1 | 245399 | 0.086531 | 98.8721 | 107.0271
Mean | [6 empty] | 0.0858 | 107.4879 | 107.4879
Median | [7 empty] | 102.0955 | 107.1995
SD | [7 empty] | 13.4843 | 0.72113
RSD | [7 empty] | 12.54% | 0.6709%
RSD Upscale Factor | [7 empty] | 0.053479
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 262b (Certified Value 104 ppb) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppb
|  |  | INAA | FA*MS | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*AAS | 0
|  |  | 0.085g | 25g | 50g | 30g | 50g | 30g | 50g | 50g | 50g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 1
| 1 | 1 | 98.8721 | 110 | 102.5082 | 106.0936 | 110 | 103.48 | 97 | 106 | 84.5 | 98 | 89 | 119 | 103 | 108 | 105 | 103 | 100.9798 | 104 | 107 | 104 | 95 | 112 | 1
| 1 | 2 | 100.5888 | 114 | 104.2813 | 105.5395 | 100 | 98.92 | 100 | 113 | 90 | 107 | 87 | 121 | 106 | 113 | 104 | 100 | 98 | 104 | 102 | 94 | 99 | 107 | 0
| 1 | 3 | 102.8113 | 104 | 106.0544 | 106.7298 | 110 | 103.08 | 100 | 116 | 86 | 111 | 88 | 119.8611 | 99 | 109 | 105 | 115 | 106 | 102 | 102 | 99 | 101 | 110 | 16
| 1 | 4 | 106.0404 | 106 | 107.8275 | 104.1008 | 100 | 102.52 | 51.7 | 108 | 87 | 99 | 93 | 121 | 97 | 111 | 105 | 109 | 103.9169 | 102 | 119 | 105 | 101 | 105 | 104.108
| 1 | 5 | 95.5319 | 109 | 103.3948 | 108.3222 | 140 | 99.1 | 97.9 | 112 | 90 | 105 | 88 | 123 | 107 | 111 | 104 | 114 | 111.9328 | 102 | 100 | 96 | 100 | 109 | 7
| 1 | 6 | 100.9293 | 107 | 106.0544 | 110.7237 | 120 | 105.55 | 100 | 110 | 87 | 104 | 92 | 119.1667 | 96 | 107 | 103 | 110 | 109.5219 | 106 | 103 | 102 | 103 | 113
|  | 7 | 2.00: 113.0539
|  | 8 | 2.00: 94.5215
|  | 9 | 2.00: 124.6022
|  | 10 | 2.00: 101.4569
|  | 11 | 2.00: 102.734
|  | 12 | 2.00: 98.0939
|  | 13 | 2.00: 141.4018
|  | 14 | 2.00: 99.4465
|  | 15 | 2.00: 120.185
|  | 16 | 2.00: 97.8839
|  | 17 | 2.00: 109.7681
|  | 18 | 2.00: 115.088
|  | 19 | 2.00: 134.3882
|  | 20 | 2.00: 92.3597
| Mean |  | 107.4879 | 108.3333 | 105.0201 | 106.9183 | 113.3333 | 102.1083 | 91.1 | 110.8333 | 87.4167 | 104 | 89.5 | 120.5046 | 101.3333 | 109.8333 | 104.3333 | 108.5 | 105.0586 | 103.3333 | 105.5 | 100 | 99.8333 | 109.3333
| Median |  | 102.0955 | 108 | 105.1679 | 106.4117 | 110 | 102.8 | 98.95 | 111 | 87 | 104.5 | 88.5 | 120.4306 | 101 | 110 | 104.5 | 109.5 | 104.9584 | 103 | 102.5 | 100.5 | 100.5 | 109.5
| Std Dev. |  | 13.4843 | 3.5024 | 1.9758 | 2.3241 | 15.0555 | 2.6098 | 19.3445 | 3.6009 | 2.2004 | 4.899 | 2.429 | 1.4957 | 4.6762 | 2.2286 | 0.8165 | 5.9582 | 5.212 | 1.633 | 7.0071 | 4.4272 | 2.7142 | 3.0111
| Rel.Std.Dev. |  | 12.54% | 3.233% | 1.881% | 2.174% | 13.28% | 2.556% | 21.23% | 3.249% | 2.517% | 4.711% | 2.714% | 1.241% | 4.615% | 2.029% | 0.7826% | 5.491% | 4.961% | 1.58% | 6.642% | 4.427% | 2.719% | 2.754%
| PDM3 |  | 3.247% | 4.059% | 0.8761% | 2.699% | 8.861% | -1.921% | -12.49% | 6.46% | -16.03% | -0.1037% | -14.03% | 15.75% | -2.665% | 5.499% | 0.2164% | 4.219% | 0.9131% | -0.7441% | 1.337% | -3.946% | -4.106% | 5.019%
| Z-Score (Absolute) |  | NA | 0.61 | 0 | 0.35 | 1.52 | 0.53 | 2.55 | 1.06 | 3.22 | 0.19 | 2.84 | 2.83 | 0.67 | 0.88 | 0.13 | 0.64 | 0.01 | 0.31 | 0.09 | 0.92 | 0.95 | 0.79
| Analytical results for Pd in OREAS 262b (Indicative Value < 5 ppb) | [23 empty] | Indicative
Pd | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.22 | ppb
|  |  | 2.00: FA*AAS | 2
|  |  | 2.00: 30g | 2
| 1 | 1 | 2.00: < 5 | 1
| 1 | 2 | 2.00: < 5 | 1
| 1 | 3 | 2.00: < 5 | 16
| 1 | 4 | 2.00: < 5 | < 5
| 1 | 5 | 2.00: < 5 | 7
| 1 | 6 | 2.00: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.00
| Analytical results for Pt in OREAS 262b (Indicative Value < 10 ppb) | [23 empty] | Indicative
Pt | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.22 | ppb
|  |  | 2.00: FA*AAS | 1
|  |  | 2.00: 30g | 1
| 1 | 1 | 2.00: < 10 | 1
| 1 | 2 | 2.00: < 10 | 1
| 1 | 3 | 2.00: < 10 | 16
| 1 | 4 | 2.00: < 10 | < 10
| 1 | 5 | 2.00: < 10 | 7
| 1 | 6 | 2.00: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 262b (Certified Value 93.5 ppb) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | ppb
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | 1
|  |  | 0.085g | 25g | 30g | 50g | 30g | 50g | 30g | 30g | 25g | 15g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 2
| 1 | 1 | 98.8721 | 74 | 112.4619 | 100 | 100 | 80 | 66.9 | 113 | 92 | 115.5 | 88 | 110 | 102.9734 | 93 | 82 | 93 | 100 | 1
| 1 | 2 | 100.5888 | 73 | 114.4071 | 80 | 100 | 80 | 80.2 | 111 | 95 | 105 | 87 | 87 | 96.9688 | 95 | 82 | 87 | 90 | 0
| 1 | 3 | 102.8113 | 69 | 116.3524 | 60 | 100 | 80 | 106 | 111 | 98 | 94.5 | 90 | 98 | 101.5403 | 103 | 85 | 104 | 130 | 16
| 1 | 4 | 106.0404 | 73 | 118.2977 | 80 | 100 | 80 | 79.3 | 112 | 96 | 105 | 87 | 90 | 99.9615 | 98 | 89 | 94 | 100 | 93.5224
| 1 | 5 | 95.5319 | 70 | 113.4345 | 80 | 100 | 80 | 81.2 | 110 | 88 | 94.5 | 89 | 87 | 103.3729 | 107 | 83 | 99 | 110 | 9
| 1 | 6 | 100.9293 | 67 | 116.3524 | 60 | 100 | 80 | 94.2 | 115 | 94 | 115.5 | 87 | 89 | 97.2317 | 91 | 84 | 88 | 100
|  | 7 | 2.00: 113.0539
|  | 8 | 2.00: 94.5215
|  | 9 | 2.00: 124.6022
|  | 10 | 2.00: 101.4569
|  | 11 | 2.00: 102.734
|  | 12 | 2.00: 98.0939
|  | 13 | 2.00: 141.4018
|  | 14 | 2.00: 99.4465
|  | 15 | 2.00: 120.185
|  | 16 | 2.00: 97.8839
|  | 17 | 2.00: 109.7681
|  | 18 | 2.00: 115.088
|  | 19 | 2.00: 134.3882
|  | 20 | 2.00: 92.3597
| Mean |  | 107.4879 | 71 | 115.2177 | 76.6667 | 100 | 80 | 84.6333 | 112 | 93.8333 | 105 | 88 | 93.5 | 100.3414 | 97.8333 | 84.1667 | 94.1667 | 105
| Median |  | 102.0955 | 71.5 | 115.3798 | 80 | 100 | 80 | 80.7 | 111.5 | 94.5 | 105 | 87.5 | 89.5 | 100.7509 | 96.5 | 83.5 | 93.5 | 100
| Std Dev. |  | 13.4843 | 2.7568 | 2.1676 | 15.0555 | 0 | 0 | 13.5824 | 1.7889 | 3.4881 | 9.3915 | 1.2649 | 9.0499 | 2.7837 | 6.1455 | 2.6394 | 6.4936 | 13.784
| Rel.Std.Dev. |  | 12.54% | 3.883% | 1.881% | 19.64% | 0% | 0% | 16.05% | 1.597% | 3.717% | 8.944% | 1.437% | 9.679% | 2.774% | 6.282% | 3.136% | 6.896% | 13.13%
| PDM3 |  | 14.93% | -24.08% | 23.2% | -18.02% | 6.926% | -14.46% | -9.505% | 19.76% | 0.3324% | 12.27% | -5.905% | -0.024% | 7.291% | 4.609% | -10% | 0.6888% | 12.27%
| Z-Score (Absolute) |  | NA | 1.62 | 1.49 | 1.22 | 0.42 | 0.98 | 0.66 | 1.26 | 0.01 | 0.77 | 0.42 | 0.04 | 0.45 | 0.27 | 0.69 | 0.01 | 0.77
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 262b (Certified Value 82.1 ppb) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | ppb
|  |  | INAA | CNL*MS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*OES | CNL*MS | CNL*AAS | CNL*AAS | 1
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 05g | 50g | 50g | 200g | 50g | 30g | 30g | 30g | 25g | 20g | 05g | 2
| 1 | 1 | 98.8721 | 110 | 97.7139 | 67.3082 | 60 | 82.4 | < 100 | 65.4122 | 74.7 | 80 | 103 | 67.68 | 70 | 85 | 79.27 | 97 | 80 | 1
| 1 | 2 | 100.5888 | 80 | 99.4041 | 64.7411 | 60 | 105.1 | < 100 | 80.7346 | 74.7 | 80 | 103 | 62.6 | 70 | 87 | 78.85 | 81.9181 | 81 | 0
| 1 | 3 | 102.8113 | 80 | 101.0943 | 65.4243 | 50 | 115.1 | 100 | 55.4609 | 73.1 | 80 | 108 | 68.01 | 70 | 85 | 78.07 | 97.9021 | 82 | 16
| 1 | 4 | 106.0404 | 80 | 102.7844 | 64.8046 | 60 | 87.7 | < 100 | 80.696 | 75.6 | 80 | 103 | 68.82 | 70 | 87 | 79.89 | 87.1089 | 76 | 82.1072
| 1 | 5 | 95.5319 | 90 | 98.559 | 68.6802 | 60 | 91.1 | < 100 | 68.0343 | 81.7 | 80 | 103 | 66.62 | 80 | 86 | 77.26 | 92.9071 | 76 | 11
| 1 | 6 | 100.9293 | 90 | 101.0943 | 69.7619 | 60 | 115.6 | 100 | 70 | 71.6 | 80 | 103 | 67 | 60 | 90 | 79.28 | 79 | 85
|  | 7 | 2.00: 113.0539
|  | 8 | 2.00: 94.5215
|  | 9 | 2.00: 124.6022
|  | 10 | 2.00: 101.4569
|  | 11 | 2.00: 102.734
|  | 12 | 2.00: 98.0939
|  | 13 | 2.00: 141.4018
|  | 14 | 2.00: 99.4465
|  | 15 | 2.00: 120.185
|  | 16 | 2.00: 97.8839
|  | 17 | 2.00: 109.7681
|  | 18 | 2.00: 115.088
|  | 19 | 2.00: 134.3882
|  | 20 | 2.00: 92.3597
| Mean |  | 107.4879 | 88.3333 | 100.1083 | 66.7867 | 58.3333 | 99.5 | 100 | 70.0563 | 75.2333 | 80 | 103.8333 | 66.7883 | 70 | 86.6667 | 78.77 | 89.306 | 80
| Median |  | 102.0955 | 85 | 100.2492 | 66.3662 | 60 | 98.1 | 100 | 69.0171 | 74.7 | 80 | 103 | 67.34 | 70 | 86.5 | 79.06 | 90.008 | 80.5
| Std Dev. |  | 13.4843 | 11.6905 | 1.8834 | 2.1297 | 4.0825 | 14.3954 | 0 | 9.6531 | 3.472 | 0 | 2.0412 | 2.1923 | 6.3246 | 1.8619 | 0.953 | 7.8953 | 3.5214
| Rel.Std.Dev. |  | 12.54% | 13.23% | 1.881% | 3.189% | 6.999% | 14.47% | 0% | 13.78% | 4.615% | 0% | 1.966% | 3.282% | 9.035% | 2.148% | 1.21% | 8.841% | 4.402%
| PDM3 |  | 30.91% | 7.583% | 21.92% | -18.66% | -28.95% | 21.18% | 21.79% | -14.68% | -8.372% | -2.566% | 26.46% | -18.66% | -14.75% | 5.553% | -4.064% | 8.768% | -2.566%
| Z-Score (Absolute) |  | NA | 0.63 | 1.45 | 0.88 | 1.47 | 1.41 | 0.89 | 0.65 | 0.29 | 0.04 | 1.71 | 0.88 | 0.66 | 0.51 | 0.04 | 0.69 | 0.04
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 262b (Certified Value 108 ppb) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | ppb
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 0
|  |  | 0.085g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 460±20 g | 1
| 1 | 1 | 98.8721 | 135 | 99 | 110 | 140 | 130 | 70 | 90 | 85 | 120 | 84 | 125 | 1
| 1 | 2 | 100.5888 | 136 | 99 | 70 | 80 | 90 | 130 | 70 | 141 | 90 | 94 | 91 | 0
| 1 | 3 | 102.8113 | 77 | 131 | 120 | 140 | 90 | 120 | 110 | 119 | 101 | 114 | 124 | 16
| 1 | 4 | 106.0404 | 81 | 99 | 130 | 140 | 150 | 120 | 90 | 66 | 132 | 60 | 103 | 108.0455
| 1 | 5 | 95.5319 | 116 | 153 | 100 | 110 | 90 | 90 | 150 | 106 | 94 | 67 | 116 | 13
| 1 | 6 | 100.9293 | 111 | 133 | 100 | 120 | 100 | 100 | 130 | 78 | 135 | 115 | 121
|  | 7 | 2.00: 113.0539
|  | 8 | 2.00: 94.5215
|  | 9 | 2.00: 124.6022
|  | 10 | 2.00: 101.4569
|  | 11 | 2.00: 102.734
|  | 12 | 2.00: 98.0939
|  | 13 | 2.00: 141.4018
|  | 14 | 2.00: 99.4465
|  | 15 | 2.00: 120.185
|  | 16 | 2.00: 97.8839
|  | 17 | 2.00: 109.7681
|  | 18 | 2.00: 115.088
|  | 19 | 2.00: 134.3882
|  | 20 | 2.00: 92.3597
| Mean |  | 107.4879 | 109.3333 | 119 | 105 | 121.6667 | 108.3333 | 105 | 106.6667 | 99.1667 | 112 | 89 | 113.3333
| Median |  | 102.0955 | 113.5 | 115 | 105 | 130 | 95 | 110 | 100 | 95.5 | 110.5 | 89 | 118.5
| Std Dev. |  | 13.4843 | 25.5552 | 23.2207 | 20.7364 | 24.0139 | 25.6255 | 22.5832 | 29.4392 | 28.0672 | 19.6061 | 23.1344 | 13.5745
| Rel.Std.Dev. |  | 12.54% | 23.37% | 19.51% | 19.75% | 19.74% | 23.65% | 21.51% | 27.6% | 28.3% | 17.51% | 25.99% | 11.98%
| PDM3 |  | -0.5161% | 1.192% | 10.14% | -2.819% | 12.61% | 0.2664% | -2.819% | -1.276% | -8.218% | 3.66% | -17.63% | 4.894%
| Z-Score (Absolute) |  | NA | 0.18 | 1.96 | 0.61 | 2.45 | 0 | 0.61 | 0.31 | 1.69 | 0.67 | 3.56 | 0.92
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 262b (Certified Value 0.158 ppm) | [19 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.16 | < 0.2 | 0.3768 | 0.2 | < 0.3 | 0.1 | < 0.5 | 0.19 | 0.2 | < 0.3 | 0.17 | < 0.1 | 0.3 | 0.16 | 0.16 | 0.16 |  |  | 1
| 1 | 2 | 0.16 | < 0.2 | < 0.1 | 0.2 | < 0.3 | 0.08 | < 0.5 | 0.18 | 0.2 | < 0.3 | 0.16 | < 0.1 | 0.3 | 0.14 | 0.15 | 0.16 |  |  | 21
| 1 | 3 | 0.22 | < 0.2 | < 0.1 | 0.2 | < 0.3 | 0.08 | < 0.5 | 0.19 | 0.1 | < 0.3 | 0.21 | < 0.1 | 0.3 | 0.14 | 0.16 | 0.13 |  |  | 16
| 1 | 4 | 0.15 | < 0.2 | < 0.1 | 0.2 | < 0.3 | 0.09 | < 0.5 | 0.12 | 0.1 | < 0.3 | 0.2 | < 0.1 | 0.3 | 0.14 | 0.15 | 0.13 |  |  | 0.15767
| 1 | 5 | 0.16 | < 0.2 | < 0.1 | 0.2 | < 0.3 | 0.07 | < 0.5 | 0.14 | 0.1 | < 0.3 | 0.17 | < 0.1 | 0.3 | 0.15 | 0.14 | 0.15 |  |  | 15
| 1 | 6 | 0.15 | < 0.2 | < 0.1 | 0.2 | < 0.3 | 0.08 | < 0.5 | 0.17 | 0.5 | < 0.3 | 0.17 | < 0.1 | 0.3 | 0.15 | 0.15 | 0.15
| Mean |  | 0.16667 |  | 0.3768 | 0.2 |  | 0.083333 |  | 0.165 | 0.2 |  | 0.18 |  | 0.3 | 0.14667 | 0.15167 | 0.14667
| Median |  | 0.16 |  | 0.3768 | 0.2 |  | 0.08 |  | 0.175 | 0.15 |  | 0.17 |  | 0.3 | 0.145 | 0.15 | 0.15
| Std Dev. |  | 0.026583 |  |  | 3.0405e-17 |  | 0.010328 |  | 0.02881 | 0.15492 |  | 0.02 |  | 0 | 0.008165 | 0.0075277 | 0.013663
| Rel.Std.Dev. |  | 15.95% |  |  | 1.52e-14% |  | 12.39% |  | 17.46% | 77.46% |  | 11.11% |  | 0% | 5.567% | 4.963% | 9.315%
| PDM3 |  | 5.708% |  | 139% | 26.85% |  | -47.15% |  | 4.651% | 26.85% |  | 14.16% |  | 90.27% | -6.977% | -3.805% | -6.977%
| Z-Score (Absolute) |  | 0.67 | 2.02 | 1.84 | NA | 0 | 2.7 | 4.05 | 0.61 | NA | 0 | 1.21 | 4.05 | NA | 0.13 | 0.07 | 0.13
Results from laboratories 1.07, 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 262b (Certified Value 5.01 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.1385 | 4.9002 | 5.3219 | 4.7297 | 4.88 | 5.43 | 5.37 | 5 | 4.87 | 5.36 | 4.682 | 5.08 | 4.9888 | 5.12 | 4.87 | 5.06 | 4.91 | 5.28 | 1
| 1 | 2 | 5.1116 | 4.9656 | 5.2816 | 4.7296 | 4.88 | 5.5 | 5.31 | 5.06 | 4.87 | 5.32 | 4.6757 | 4.96 | 4.9961 | 5.09 | 4.99 | 4.79 | 4.81 | 5.25 | 0
| 1 | 3 | 4.9954 | 4.9013 | 5.1823 | 4.6435 | 4.89 | 5.32 | 5.29 | 5.05 | 4.83 | 5.29 | 4.7437 | 4.99 | 4.8698 | 5.01 | 4.88 | 4.97 | 4.92 | 5.35 | 16
| 1 | 4 | 5.1619 | 4.9478 | 5.1399 | 4.6887 | 4.91 | 5.45 | 5.41 | 5.05 | 4.84 | 5.01 | 4.6157 | 5.18 | 4.8553 | 4.97 | 5.01 | 4.88 | 4.75 | 5.3 | 5.0095
| 1 | 5 | 5.1501 | 4.8814 | 5.2475 | 4.71 | 4.9 | 5.45 | 5.36 | 4.91 | 4.81 | 4.97 | 4.8101 | 4.97 | 4.8056 | 4.99 | 4.94 | 4.96 | 4.8 | 5.08 | 16
| 1 | 6 | 4.9707 | 4.9298 | 5.2501 | 4.6415 | 4.87 | 5.41 | 5.31 | 5.07 | 4.75 | 5.13 | 4.7756 | 4.96 | 4.9921 | 5.07 | 4.93 | 4.78 | 4.72 | 5.03
| Mean |  | 5.088 | 4.921 | 5.2372 | 4.6905 | 4.8883 | 5.4267 | 5.3417 | 5.0233 | 4.8283 | 5.18 | 4.7171 | 5.0233 | 4.9179 | 5.0417 | 4.9367 | 4.9067 | 4.8183 | 5.215
| Median |  | 5.125 | 4.9155 | 5.2488 | 4.6993 | 4.885 | 5.44 | 5.335 | 5.05 | 4.835 | 5.21 | 4.7129 | 4.98 | 4.9293 | 5.04 | 4.935 | 4.92 | 4.805 | 5.265
| Std Dev. |  | 0.083377 | 0.032188 | 0.066164 | 0.04016 | 0.01472 | 0.060222 | 0.045789 | 0.060553 | 0.044907 | 0.16709 | 0.07212 | 0.089144 | 0.084249 | 0.060139 | 0.056451 | 0.11021 | 0.081833 | 0.12911
| Rel.Std.Dev. |  | 1.639% | 0.6541% | 1.263% | 0.8562% | 0.3011% | 1.11% | 0.8572% | 1.205% | 0.9301% | 3.226% | 1.529% | 1.775% | 1.713% | 1.193% | 1.143% | 2.246% | 1.698% | 2.476%
| PDM3 |  | 1.568% | -1.766% | 4.546% | -6.368% | -2.418% | 8.328% | 6.631% | 0.2767% | -3.616% | 3.404% | -5.836% | 0.2767% | -1.827% | 0.6427% | -1.453% | -2.052% | -3.816% | 4.103%
| Z-Score (Absolute) |  | 0.56 | 0.31 | 1.34 | 1.5 | 0.48 | 2.32 | 1.88 | 0.23 | 0.79 | 1.04 | 1.37 | 0.23 | 0.32 | 0.32 | 0.23 | 0.38 | 0.84 | 1.22
| Analytical results for As in OREAS 262b (Certified Value 30.3 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 32.7 | 37.5199 | 142.4188 | 30 | 27.7 | 26.8 | 27.8 | 31.4 | 30 | 15.4324 | 29.4 | 30.9072 | 29.9 | 31 | 32.3 | 32.7 | 24 |  | 1
| 1 | 2 | 32.8 | 35.878 | 101.2439 | 30 | 30.4 | 26.1 | 27.9533 | 31.1 | 29 | 15.2105 | 30.4 | 31.1025 | 29.3 | 31.7 | 32.2 | 31.9 | 23 |  | 22
| 1 | 3 | 32.8 | 37.5256 | 152.2941 | 31 | 30.3 | 26.4 | 28.195 | 29.4 | 31 | 15.5318 | 29.7 | 29.3724 | 28.9 | 31.9 | 32.5 | 32.5 | 26 |  | 16
| 1 | 4 | 33.5 | 36.4907 | 100.2193 | 30 | 28.2 | 26.4 | 29.3067 | 31.9 | 31 | 15.6095 | 32.2 | 30.0596 | 29 | 31.5 | 32.5 | 32.5 | 24 |  | 30.2542
| 1 | 5 | 32.6 | 37.0975 | 114.9827 | 29 | 28.5 | 25.3 | 27.5033 | 30.8 | 30 | 15.5131 | 31.7 | 28.9531 | 29.1 | 31.5 | 31.5 | 34.2 | 24 |  | 17
| 1 | 6 | 32 | 36.4357 | 124.1062 | 29 | 28.1 | 25.1 | 27.8233 | 30.7 | 29 | 15.3449 | 30.1 | 29.6082 | 30.2 | 31 | 33 | 31.5 | 26
| Mean |  | 32.7333 | 36.8246 | 122.5442 | 29.8333 | 28.8667 | 26.0167 | 28.0969 | 30.8833 | 30 | 15.4404 | 30.5833 | 30.0005 | 29.4 | 31.4333 | 32.3333 | 32.55 | 24.5
| Median |  | 32.75 | 36.7941 | 119.5444 | 30 | 28.35 | 26.25 | 27.8883 | 30.95 | 30 | 15.4727 | 30.25 | 29.8339 | 29.2 | 31.5 | 32.4 | 32.5 | 24
| Std Dev. |  | 0.48028 | 0.66454 | 21.4092 | 0.75277 | 1.1776 | 0.67355 | 0.63389 | 0.84715 | 0.89443 | 0.14441 | 1.1232 | 0.85854 | 0.52915 | 0.36697 | 0.49261 | 0.92466 | 1.2247
| Rel.Std.Dev. |  | 1.467% | 1.805% | 17.47% | 2.523% | 4.079% | 2.589% | 2.256% | 2.743% | 2.981% | 0.9353% | 3.673% | 2.862% | 1.8% | 1.167% | 1.524% | 2.841% | 4.999%
| PDM3 |  | 8.194% | 21.72% | 305% | -1.391% | -4.586% | -14.01% | -7.131% | 2.079% | -0.8403% | -48.96% | 1.088% | -0.8387% | -2.823% | 3.897% | 6.872% | 7.588% | -19.02%
| Z-Score (Absolute) |  | 0.97 | 2.42 | 32.78 | 0.06 | 0.4 | 1.41 | 0.67 | 0.31 | 0 | 5.16 | 0.21 | 0 | 0.21 | 0.51 | 0.83 | 0.9 | 1.95
| Analytical results for Ba in OREAS 262b (Certified Value 336 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 345.3 | 335.918 | 332.3218 | 325.9761 | 328 | 293 | 344 | 350 | 330 | 221 | 318.5862 | 350 | 341.3617 | 336 | 340 | 360 | 330 | 368 | 1
| 1 | 2 | 352.4 | 340.604 | 323.0574 | 324.8427 | 332 | 324 | 346 | 349 | 330 | 356 | 318.7718 | 340 | 340.7176 | 332 | 340 | 340 | 320 | 344 | 23
| 1 | 3 | 351 | 336.699 | 327.4308 | 318.3685 | 328 | 336 | 345 | 348 | 330 | 229 | 318.7308 | 340 | 334.6473 | 324 | 340 | 350 | 330 | 357 | 16
| 1 | 4 | 345.4 | 340.219 | 327.9308 | 323.2905 | 334 | 329 | 349 | 343 | 330 | 221 | 319.437 | 350 | 334.2642 | 329 | 340 | 350 | 320 | 364 | 336.4853
| 1 | 5 | 344.3 | 337.15 | 323.1837 | 323.9243 | 330 | 348 | 343 | 342 | 330 | 218 | 318.9925 | 340 | 331.1104 | 327 | 350 | 350 | 330 | 357 | 18
| 1 | 6 | 346.4 | 338.393 | 330.0243 | 319.0298 | 328 | 326 | 345 | 351 | 330 | 223 | 318.5299 | 330 | 336.5918 | 335 | 340 | 340 | 320 | 353
| Mean |  | 347.4667 | 338.1638 | 327.3248 | 322.572 | 330 | 326 | 345.3333 | 347.1667 | 330 | 244.6667 | 318.8414 | 341.6667 | 336.4488 | 330.5 | 341.6667 | 348.3333 | 325 | 357.1667
| Median |  | 345.9 | 337.7715 | 327.6808 | 323.6074 | 329 | 327.5 | 345 | 348.5 | 330 | 222 | 318.7513 | 340 | 335.6196 | 330.5 | 340 | 350 | 325 | 357
| Std Dev. |  | 3.375 | 1.9208 | 3.6877 | 3.1402 | 2.5298 | 18.363 | 2.0656 | 3.7639 | 0 | 54.665 | 0.33371 | 7.5277 | 3.9723 | 4.6797 | 4.0825 | 7.5277 | 5.4772 | 8.4242
| Rel.Std.Dev. |  | 0.9713% | 0.568% | 1.127% | 0.9735% | 0.7666% | 5.633% | 0.5981% | 1.084% | 0% | 22.34% | 0.1047% | 2.203% | 1.181% | 1.416% | 1.195% | 2.161% | 1.685% | 2.359%
| PDM3 |  | 3.264% | 0.4988% | -2.722% | -4.135% | -1.927% | -3.116% | 2.63% | 3.174% | -1.927% | -27.29% | -5.244% | 1.54% | -0.01085% | -1.779% | 1.54% | 3.521% | -3.413% | 6.146%
| Z-Score (Absolute) |  | 1.13 | 0.38 | 0.5 | 0.88 | 0.28 | 0.6 | 0.96 | 1.11 | 0.28 | 7.19 | 1.18 | 0.66 | 0.24 | 0.24 | 0.66 | 1.2 | 0.69 | 1.92
| Analytical results for Be in OREAS 262b (Certified Value 1.52 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.52 | 1.5823 | 1.1695 | 1.4 | 1.8 | 1.54 | < 5 | 1.5 | 2 | 1.264 | 1.55 | 1.5103 | 1.5 | 1.48 | 1.58 | 1.44 | 1.65 |  | 1
| 1 | 2 | 1.51 | 1.5544 | 1.1865 | 1.4 | 1.5 | 1.53 | < 5 | 1.48 | 2 | 1.2525 | 1.52 | 1.5854 | 1.6 | 1.52 | 1.48 | 1.42 | 1.61 |  | 24
| 1 | 3 | 1.52 | 1.5088 | 1.08 | 1.4 | 1.5 | 1.52 | < 5 | 1.48 | 2 | 1.2766 | 1.53 | 1.5358 | 1.5 | 1.5 | 1.52 | 1.43 | 1.73 |  | 16
| 1 | 4 | 1.54 | 1.5425 | 1.1793 | 1.4 | 1.5 | 1.58 | < 5 | 1.48 | 2 | 1.1962 | 1.55 | 1.593 | 1.5 | 1.54 | 1.51 | 1.4 | 1.83 |  | 1.5178
| 1 | 5 | 1.5 | 1.5225 | 1.1669 | 1.4 | 1.5 | 1.52 | < 5 | 1.48 | 1 | 1.2152 | 1.51 | 1.4554 | 1.5 | 1.5 | 1.51 | 1.41 | 1.57 |  | 19
| 1 | 6 | 1.5 | 1.5699 | 1.1324 | 1.4 | 1.4 | 1.53 | < 5 | 1.46 | 2 | 1.2888 | 1.52 | 1.4726 | 1.5 | 1.5 | 1.47 | 1.39 | 1.66
| Mean |  | 1.515 | 1.5467 | 1.1524 | 1.4 | 1.5333 | 1.5367 |  | 1.48 | 1.8333 | 1.2489 | 1.53 | 1.5254 | 1.5167 | 1.5067 | 1.5117 | 1.415 | 1.675
| Median |  | 1.515 | 1.5484 | 1.1682 | 1.4 | 1.5 | 1.53 |  | 1.48 | 2 | 1.2582 | 1.525 | 1.5231 | 1.5 | 1.5 | 1.51 | 1.415 | 1.655
| Std Dev. |  | 0.015166 | 0.027958 | 0.040077 | 2.4324e-16 | 0.13663 | 0.022509 |  | 0.012649 | 0.40825 | 0.036089 | 0.016733 | 0.056909 | 0.040825 | 0.020656 | 0.038687 | 0.018708 | 0.092898
| Rel.Std.Dev. |  | 1.001% | 1.808% | 3.478% | 1.737e-14% | 8.91% | 1.465% |  | 0.8547% | 22.27% | 2.89% | 1.094% | 3.731% | 2.692% | 1.371% | 2.559% | 1.322% | 5.546%
| PDM3 |  | -0.1841% | 1.906% | -24.07% | -7.761% | 1.024% | 1.243% |  | -2.49% | 20.79% | -17.72% | 0.8042% | 0.5031% | -0.07429% | -0.7331% | -0.4037% | -6.773% | 10.36%
| Z-Score (Absolute) |  | 0.02 | 0.81 | 9.48 | NA | NA | 0.54 | 25.66 | 0.93 | NA | 6.96 | 0.37 | 0.25 | 0.02 | 0.24 | 0.11 | 2.63 | 4.15
Results from laboratories 1.07 and 1.09 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 262b (Certified Value 0.73 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 0.78 | 0.51135 | 0.6728 | 2.732 | 0.74 | 0.81 | 0.7 | < 5 | 0.79 | 0.7 | 0.7574 | 0.66 | 0.72343 | 0.76 | 0.73 | 0.7 | 0.73 | 0.86 | 1
| 1 | 2 | 0.7 | 0.57041 | 0.71369 | 3.3764 | 0.74 | 0.76 | 0.68 | < 5 | 0.81 | 0.8 | 0.7847 | 0.68 | 0.7397 | 0.73 | 0.69 | 0.69 | 0.72 | 0.81 | 25
| 1 | 3 | 0.79 | 0.5515 | 0.6816 | 3.8988 | 0.73 | 0.75 | 0.67 | < 5 | 0.87 | 0.7 | 0.7524 | 0.67 | 0.68648 | 0.72 | 0.7 | 0.68 | 0.73 | 0.87 | 16
| 1 | 4 | 0.7 | 0.45052 | 0.68303 | 1.8786 | 0.76 | 0.73 | 0.71 | < 5 | 0.78 | 0.7 | 0.7535 | 0.72 | 0.73136 | 0.73 | 0.7 | 0.69 | 0.75 | 0.85 | 0.73359
| 1 | 5 | 0.71 | 0.4656 | 0.66446 | 4.1718 | 0.75 | 0.73 | 0.67 | < 5 | 0.76 | 0.7 | 0.7584 | 0.72 | 0.67773 | 0.74 | 0.75 | 0.7 | 0.74 | 0.86 | 20
| 1 | 6 | 0.74 | 0.49103 | 0.67618 | 3.72 | 0.77 | 0.74 | 0.67 | < 5 | 0.81 | 0.8 | 0.7657 | 0.68 | 0.67263 | 0.75 | 0.73 | 0.7 | 0.73 | 0.84
| Mean |  | 0.73667 | 0.50674 | 0.68196 | 3.2963 | 0.74833 | 0.75333 | 0.68333 |  | 0.80333 | 0.73333 | 0.76202 | 0.68833 | 0.70522 | 0.73833 | 0.71667 | 0.69333 | 0.73333 | 0.84833
| Median |  | 0.725 | 0.50119 | 0.67889 | 3.5482 | 0.745 | 0.745 | 0.675 |  | 0.8 | 0.7 | 0.7579 | 0.68 | 0.70495 | 0.735 | 0.715 | 0.695 | 0.73 | 0.855
| Std Dev. |  | 0.040332 | 0.047282 | 0.016921 | 0.85367 | 0.01472 | 0.030111 | 0.017512 |  | 0.037771 | 0.05164 | 0.012062 | 0.025626 | 0.029575 | 0.01472 | 0.023381 | 0.008165 | 0.010328 | 0.02137
| Rel.Std.Dev. |  | 5.475% | 9.331% | 2.481% | 25.9% | 1.967% | 3.997% | 2.563% |  | 4.702% | 7.042% | 1.583% | 3.723% | 4.194% | 1.994% | 3.262% | 1.178% | 1.408% | 2.519%
| PDM3 |  | 0.4189% | -30.92% | -7.039% | 349.3% | 2.009% | 2.691% | -6.851% |  | 9.507% | -0.03549% | 3.874% | -6.17% | -3.868% | 0.6461% | -2.307% | -5.488% | -0.03549% | 15.64%
| Z-Score (Absolute) |  | 0 | 3.58 | 0.85 | 39.83 | 0.18 | 0.26 | 0.83 | 27.44 | 1.04 | NA | 0.39 | 0.75 | 0.49 | 0.03 | 0.31 | 0.67 | 0.05 | 1.74
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 262b (Certified Value 0.384 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.389 | 0.37829 | 0.39839 | 1.2179 | 0.371 | 0.37 | 0.37 | 0.37 | 0.38 | 0.44 | 0.39778 | 0.42 | 0.38347 | 0.38 | 0.4 | 0.44 | 0.39 | 0.393 | 1
| 1 | 2 | 0.3882 | 0.38433 | 0.38667 | 1.1139 | 0.374 | 0.37 | 0.37 | 0.38 | 0.38 | 0.41 | 0.39318 | 0.41 | 0.38928 | 0.38 | 0.4 | 0.39 | 0.38 | 0.379 | 0
| 1 | 3 | 0.3784 | 0.37624 | 0.38635 | 1.0927 | 0.371 | 0.36 | 0.37 | 0.39 | 0.38 | 0.45 | 0.39488 | 0.41 | 0.38316 | 0.37 | 0.39 | 0.4 | 0.38 | 0.376 | 16
| 1 | 4 | 0.3883 | 0.38296 | 0.39169 | 1.1133 | 0.373 | 0.35 | 0.36 | 0.38 | 0.38 | 0.43 | 0.39779 | 0.41 | 0.38294 | 0.36 | 0.4 | 0.4 | 0.37 | 0.395 | 0.38385
| 1 | 5 | 0.3872 | 0.37943 | 0.38823 | 1.1134 | 0.375 | 0.35 | 0.37 | 0.37 | 0.38 | 0.42 | 0.40678 | 0.41 | 0.3766 | 0.37 | 0.4 | 0.41 | 0.38 | 0.364 | 21
| 1 | 6 | 0.3738 | 0.38619 | 0.39467 | 1.0923 | 0.373 | 0.37 | 0.37 | 0.39 | 0.38 | 0.44 | 0.40518 | 0.41 | 0.38847 | 0.37 | 0.39 | 0.39 | 0.37 | 0.377
| Mean |  | 0.38415 | 0.38124 | 0.391 | 1.1239 | 0.37283 | 0.36167 | 0.36833 | 0.38 | 0.38 | 0.43167 | 0.39926 | 0.41167 | 0.38399 | 0.37167 | 0.39667 | 0.405 | 0.37833 | 0.38067
| Median |  | 0.3877 | 0.3812 | 0.38996 | 1.1134 | 0.373 | 0.365 | 0.37 | 0.38 | 0.38 | 0.435 | 0.39778 | 0.41 | 0.38331 | 0.37 | 0.4 | 0.4 | 0.38 | 0.378
| Std Dev. |  | 0.0064286 | 0.0038451 | 0.0048313 | 0.047204 | 0.0016021 | 0.0098319 | 0.0040825 | 0.0089443 | 6.0809e-17 | 0.01472 | 0.0055179 | 0.0040825 | 0.0045776 | 0.0075277 | 0.005164 | 0.018708 | 0.0075277 | 0.011605
| Rel.Std.Dev. |  | 1.673% | 1.009% | 1.236% | 4.2% | 0.4297% | 2.719% | 1.108% | 2.354% | 1.6e-14% | 3.41% | 1.382% | 0.9917% | 1.192% | 2.025% | 1.302% | 4.619% | 1.99% | 3.048%
| PDM3 |  | 0.07902% | -0.6787% | 1.864% | 192.8% | -2.869% | -5.778% | -4.042% | -1.002% | -1.002% | 12.46% | 4.017% | 7.248% | 0.03654% | -3.173% | 3.34% | 5.511% | -1.436% | -0.8285%
| Z-Score (Absolute) |  | 0.1 | 0.09 | 0.55 | 48.23 | 0.64 | 1.36 | 0.93 | 0.17 | 0.17 | 3.19 | 1.08 | 1.89 | 0.09 | 0.71 | 0.91 | 1.46 | 0.28 | 0.13
| Analytical results for Cd in OREAS 262b (Certified Value 0.062 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.07 | 0.23367 | 0.1881 | 0.07 | < 0.3 | 0.26 | < 0.5 | 0.03 | 0.3 | 0.0722 | 0.04 | 0.064569 | 0.18 | 0.06 | 0.06 | 0.06 |  |  | 1
| 1 | 2 | 0.06 | 0.22652 | 0.1783 | 0.06 | < 0.3 | 0.25 | < 0.5 | 0.08 | 0.3 | 0.0755 | 0.06 | 0.060869 | 0.18 | 0.04 | 0.05 | 0.06 |  |  | 26
| 1 | 3 | 0.07 | 0.23913 | 0.3756 | 0.07 | < 0.3 | 0.25 | < 0.5 | 0.11 | 0.3 | 0.0775 | 0.05 | 0.055819 | 0.17 | 0.06 | 0.05 | 0.05 |  |  | 16
| 1 | 4 | 0.07 | 0.23823 | 0.2346 | 0.09 | < 0.3 | 0.26 | < 0.5 | 0.07 | 0.2 | 0.0861 | 0.03 | < 0.05 | 0.19 | 0.06 | 0.05 | 0.05 |  |  | 0.061625
| 1 | 5 | 0.07 | 0.24349 | 0.2149 | 0.09 | < 0.3 | 0.24 | < 0.5 | 0.02 | 0.3 | 0.0669 | 0.08 | 0.060036 | 0.17 | 0.07 | 0.05 | 0.06 |  |  | 22
| 1 | 6 | 0.06 | 0.22979 | 0.1695 | 0.08 | < 0.3 | 0.25 | < 0.5 | 0.05 | 0.4 | 0.0856 | 0.03 | < 0.05 | 0.17 | 0.05 | 0.08 | 0.06
| Mean |  | 0.066667 | 0.23514 | 0.22683 | 0.076667 |  | 0.25167 |  | 0.06 | 0.3 | 0.0773 | 0.048333 | 0.060323 | 0.17667 | 0.056667 | 0.056667 | 0.056667
| Median |  | 0.07 | 0.23595 | 0.2015 | 0.075 |  | 0.25 |  | 0.06 | 0.3 | 0.0765 | 0.045 | 0.060453 | 0.175 | 0.06 | 0.05 | 0.06
| Std Dev. |  | 0.005164 | 0.0063287 | 0.076782 | 0.012111 |  | 0.0075277 |  | 0.033466 | 0.063246 | 0.0075337 | 0.019408 | 0.0035914 | 0.008165 | 0.010328 | 0.012111 | 0.005164
| Rel.Std.Dev. |  | 7.746% | 2.691% | 33.85% | 15.8% |  | 2.991% |  | 55.78% | 21.08% | 9.746% | 40.15% | 5.954% | 4.622% | 18.23% | 21.37% | 9.113%
| PDM3 |  | 8.182% | 281.6% | 268.1% | 24.41% |  | 308.4% |  | -2.637% | 386.8% | 25.44% | -21.57% | -2.112% | 186.7% | -8.046% | -8.046% | -8.046%
| Z-Score (Absolute) |  | 0.24 | 3.8 | 3.6 | 0 | 1.76 | 4.2 | 4.16 | 0.4 | NA | 0.02 | 0.68 | 0.67 | 2.4 | 0.48 | 0.48 | 0.48
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 262b (Certified Value 68 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 67.1 | 68.3203 | 64.2434 | 66 | 54.96 | 66.3 | 68 | 34.9198 | 65.1 | 70.4 | 71.2 | 69.9 | 66 | [5 empty] | 1
| 1 | 2 | 68.25 | 68.8185 | 61.3712 | 68.6 | 53.6 | 67.3 | 73 | 39.2605 | 67.7 | 70.4 | 69.7 | 71.2 | 67.5 | [5 empty] | 27
| 1 | 3 | 66.84 | 69.1624 | 62.1678 | 69.5 | 55.65 | 66.4 | 66 | 38.2751 | 66.8 | 68.3 | 69.5 | 69.5 | 66.5 | [5 empty] | 16
| 1 | 4 | 68.22 | 68.2147 | 65.3853 | 68 | 52.17 | 66.8 | 63 | 34.7108 | 69.4 | 69.6 | 71.8 | 68.9 | 68.1 | [5 empty] | 67.7365
| 1 | 5 | 68.04 | 69.1904 | 61.6009 | 67.8 | 49.48 | 65.4 | 66 | 33.2428 | 69.6 | 69.9 | 71.4 | 69.7 | 70 | [5 empty] | 23
| 1 | 6 | 68.22 | 68.9731 | 64.491 | 67.1 | 48.93 | 66.4 | 61 | 36.3224 | 66.6 | 72.2 | 68.7 | 69.5 | 64.3
| Mean |  | 67.7783 | 68.7799 | 63.21 | 67.8333 | 52.465 | 66.4333 | 66.1667 | 36.1219 | 67.5333 | 70.1333 | 70.3833 | 69.7833 | 67.0667
| Median |  | 68.13 | 68.8958 | 63.2056 | 67.9 | 52.885 | 66.4 | 66 | 35.6211 | 67.25 | 70.15 | 70.45 | 69.6 | 67
| Std Dev. |  | 0.63587 | 0.42057 | 1.7028 | 1.2078 | 2.7982 | 0.62823 | 4.1673 | 2.2914 | 1.7386 | 1.2738 | 1.2481 | 0.7705 | 1.948
| Rel.Std.Dev. |  | 0.9382% | 0.6115% | 2.694% | 1.78% | 5.334% | 0.9456% | 6.298% | 6.344% | 2.574% | 1.816% | 1.773% | 1.104% | 2.905%
| PDM3 |  | 0.06176% | 1.54% | -6.683% | 0.143% | -22.55% | -1.924% | -2.318% | -46.67% | -0.2999% | 3.538% | 3.908% | 3.022% | -0.9889%
| Z-Score (Absolute) |  | 0.12 | 0.62 | 2.13 | 0.15 | 7.43 | 0.54 | 0.67 | 15.5 | 0 | 1.28 | 1.41 | 1.11 | 0.23
| Analytical results for Co in OREAS 262b (Certified Value 5.74 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 5.8 | 6.084 | 5.81 | 9.3296 | 5 | 6 | 6.3 | 7.0745 | 5.4 | 5.8 | 5.0718 | 5.3 | 5.6038 | 5.86 | 5.6 | 6.4 | 5.5 | 6.1 | 1
| 1 | 2 | 5.8 | 6.183 | 5.5756 | 5.8151 | 6 | 6 | 6.3 | 7.3207 | 5.4 | 5.6 | 5.0943 | 5.5 | 5.2312 | 5.79 | 5.6 | 6.6 | 5.4 | 5.9 | 28
| 1 | 3 | 5.8 | 5.796 | 5.7441 | 5.957 | 6 | 6 | 6.3 | 7.246 | 5.5 | 5.7 | 5.1422 | 5.4 | 5.6728 | 5.77 | 5.7 | 6.5 | 5.4 | 6.1 | 16
| 1 | 4 | 5.8 | 5.796 | 5.6964 | 5.8976 | 5 | 6 | 6.3 | 7.4503 | 5.5 | 5.3 | 5.0768 | 5.7 | 5.6892 | 5.68 | 5.7 | 6.5 | 5.3 | 6.3 | 5.7369
| 1 | 5 | 5.8 | 5.823 | 5.7679 | 5.6248 | 6 | 6 | 6.3 | 7.2096 | 5.4 | 5.5 | 5.2365 | 5.7 | 5.2722 | 5.68 | 5.8 | 6.2 | 5.7 | 6 | 24
| 1 | 6 | 5.6 | 6.174 | 5.7083 | 5.9253 | 5 | 6 | 6.2 | 7.3235 | 5.4 | 5.6 | 5.1995 | 5.5 | 5.4519 | 5.88 | 5.6 | 6.1 | 5.3 | 5.9
| Mean |  | 5.7667 | 5.976 | 5.7171 | 6.4249 | 5.5 | 6 | 6.2833 | 7.2708 | 5.4333 | 5.5833 | 5.1369 | 5.5167 | 5.4869 | 5.7767 | 5.6667 | 6.3833 | 5.4333 | 6.05
| Median |  | 5.8 | 5.9535 | 5.7262 | 5.9115 | 5.5 | 6 | 6.3 | 7.2834 | 5.4 | 5.6 | 5.1182 | 5.5 | 5.5278 | 5.78 | 5.65 | 6.45 | 5.4 | 6.05
| Std Dev. |  | 0.08165 | 0.19075 | 0.080556 | 1.428 | 0.54772 | 0 | 0.040825 | 0.12678 | 0.05164 | 0.17224 | 0.068611 | 0.16021 | 0.20093 | 0.085479 | 0.08165 | 0.19408 | 0.15055 | 0.15166
| Rel.Std.Dev. |  | 1.416% | 3.192% | 1.409% | 22.23% | 9.959% | 0% | 0.6497% | 1.744% | 0.9504% | 3.085% | 1.336% | 2.904% | 3.662% | 1.48% | 1.441% | 3.04% | 2.771% | 2.507%
| PDM3 |  | 0.5182% | 4.167% | -0.3465% | 11.99% | -4.13% | 4.585% | 9.524% | 26.74% | -5.292% | -2.677% | -10.46% | -3.84% | -4.359% | 0.6925% | -1.225% | 11.27% | -5.292% | 5.457%
| Z-Score (Absolute) |  | 0.06 | 0.56 | 0.06 | 1.64 | NA | NA | 1.3 | 3.66 | 0.74 | 0.38 | 1.45 | 0.54 | 0.61 | 0.08 | 0.18 | 1.54 | 0.74 | 0.74
Results from laboratories 1.07 and 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 262b (Certified Value 49.7 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
1
| 1 | 1 | 55 | 48.3927 | 95.8099 | 53 | 49 | 49 | 51.7453 | 47 | 37 | 50.4408 | 53 | 52.093 | 44 | 49 | 48 | 45 | 55 |  | 1
| 1 | 2 | 55 | 46.0593 | 81.2609 | 52 | 48 | 58 | 51.3445 | 47 | 39 | 50.34 | 52 | 54.0801 | 44 | 50 | 49 | 45 | 55 |  | 29
| 1 | 3 | 53 | 46.8389 | 79.1537 | 53 | 43 | 56 | 50.6303 | 47 | 37 | 50.9412 | 52 | 52.3115 | 45 | 48 | 48 | 46 | 58 |  | 16
| 1 | 4 | 55 | 48.3694 | 86.7639 | 54 | 46 | 60 | 50.0458 | 48 | 40 | 50.0437 | 51 | 54.0897 | 44 | 50 | 48 | 44 | 56 |  | 49.6568
| 1 | 5 | 56 | 48.1985 | 85.1495 | 55 | 46 | 60 | 53.1735 | 47 | 36 | 51.6414 | 53 | 52.6098 | 42 | 49 | 49 | 46 | 57 |  | 25
| 1 | 6 | 51 | 48.6705 | 81.5305 | 53 | 42 | 57 | 47.2184 | 46 | 49 | 51.8416 | 53 | 53.1018 | 45 | 49 | 47 | 45 | 56
| Mean |  | 54.1667 | 47.7549 | 84.9447 | 53.3333 | 45.6667 | 56.6667 | 50.693 | 47 | 39.6667 | 50.8748 | 52.3333 | 53.0477 | 44 | 49.1667 | 48.1667 | 45.1667 | 56.1667
| Median |  | 55 | 48.2839 | 83.34 | 53 | 46 | 57.5 | 50.9874 | 47 | 38 | 50.691 | 52.5 | 52.8558 | 44 | 49 | 48 | 45 | 56
| Std Dev. |  | 1.8348 | 1.052 | 6.0028 | 1.0328 | 2.7325 | 4.0825 | 2.009 | 0.63246 | 4.8028 | 0.73376 | 0.8165 | 0.87174 | 1.0954 | 0.75277 | 0.75277 | 0.75277 | 1.169
| Rel.Std.Dev. |  | 3.387% | 2.203% | 7.067% | 1.936% | 5.984% | 7.204% | 3.963% | 1.346% | 12.11% | 1.442% | 1.56% | 1.643% | 2.49% | 1.531% | 1.563% | 1.667% | 2.081%
| PDM3 |  | 9.082% | -3.83% | 71.06% | 7.404% | -8.035% | 14.12% | 2.087% | -5.35% | -20.12% | 2.453% | 5.39% | 6.829% | -11.39% | -0.9871% | -3.001% | -9.042% | 13.11%
| Z-Score (Absolute) |  | 0.67 | 0.57 | 6.65 | 0.51 | 0.98 | 1.16 | 0 | 0.72 | 2.14 | 0.04 | 0.32 | 0.46 | 1.3 | 0.3 | 0.49 | 1.07 | 1.06
| Analytical results for Cs in OREAS 262b (Certified Value 5.02 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 5.11 | 5.1441 | 4.9 | 4.73 | 5.38 | 4.97 | 4.6982 | 4.78 | 5.1499 | 4.87 | 5.23 | 5.2 | 4.97 | 5.53 |  |  |  |  | 1
| 1 | 2 | 5.09 | 4.9362 | 5 | 5.1 | 5.35 | 5.04 | 4.8701 | 4.92 | 5.1879 | 4.82 | 4.94 | 5.32 | 4.95 | 5.17 |  |  |  |  | 30
| 1 | 3 | 5.03 | 5.089 | 4.9 | 5.14 | 5.39 | 5.02 | 4.9885 | 4.91 | 5.002 | 4.71 | 4.95 | 5.16 | 4.99 | 5.53 |  |  |  |  | 16
| 1 | 4 | 5.04 | 5.1145 | 5 | 5.12 | 5.44 | 5.05 | 4.8538 | 4.91 | 5.0682 | 4.77 | 5.02 | 4.94 | 4.92 | 5.51 |  |  |  |  | 5.0194
| 1 | 5 | 4.91 | 5.1912 | 4.8 | 5.22 | 5.34 | 4.98 | 4.7485 | 5.16 | 5.0048 | 4.77 | 5.24 | 5.2 | 5.13 | 5.34 |  |  |  |  | 26
| 1 | 6 | 4.81 | 4.9973 | 4.8 | 4.91 | 5.31 | 5.03 | 5.1321 | 4.95 | 4.9444 | 4.9 | 4.87 | 5.04 | 4.84 | 5.2
| Mean |  | 4.9983 | 5.0787 | 4.9 | 5.0367 | 5.3683 | 5.015 | 4.8819 | 4.9383 | 5.0595 | 4.8067 | 5.0417 | 5.1433 | 4.9667 | 5.38
| Median |  | 5.035 | 5.1018 | 4.9 | 5.11 | 5.365 | 5.025 | 4.862 | 4.915 | 5.0365 | 4.795 | 4.985 | 5.18 | 4.96 | 5.425
| Std Dev. |  | 0.11566 | 0.095153 | 0.089443 | 0.18184 | 0.04535 | 0.032711 | 0.15902 | 0.12352 | 0.094106 | 0.070616 | 0.15715 | 0.13411 | 0.095638 | 0.16733
| Rel.Std.Dev. |  | 2.314% | 1.874% | 1.825% | 3.61% | 0.8448% | 0.6523% | 3.257% | 2.501% | 1.86% | 1.469% | 3.117% | 2.608% | 1.926% | 3.11%
| PDM3 |  | -0.4196% | 1.182% | -2.379% | 0.3442% | 6.952% | -0.08751% | -2.74% | -1.615% | 0.7995% | -4.238% | 0.4438% | 2.469% | -1.05% | 7.184%
| Z-Score (Absolute) |  | 0.25 | 0.49 | 1.16 | 0.1 | 3.15 | 0.1 | 1.32 | 0.8 | 0.31 | 2.02 | 0.15 | 1.08 | 0.54 | 3.26
| Analytical results for Cu in OREAS 262b (Certified Value 70 ppm) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 70.3 | 70.7577 | 66.4963 | 69 | 73 | 72 | 72.7045 | 66.4 | 76.6 | 72.5089 | 79.2 | 66.6129 | 69.8 | 72.7 | 72 | 69.3 | 74.5 |  | 1
| 1 | 2 | 73.1 | 66.1477 | 67.0934 | 69 | 75 | 71 | 71.5785 | 70.9 | 74.8 | 65.3 | 67.6 | 65.6596 | 68.3 | 73 | 73.8 | 68.9 | 70 |  | 31
| 1 | 3 | 74.6 | 68.9063 | 65.6836 | 69 | 73 | 70 | 70.1085 | 64.3 | 77.1 | 62.4958 | 72 | 68.3228 | 70.5 | 74.6 | 71.2 | 67.2 | 71 |  | 16
| 1 | 4 | 71.9 | 68.1935 | 67.5505 | 69 | 74 | 71 | 71.0378 | 66.4 | 76.3 | 55.0578 | 68.3 | 67.1782 | 67.1 | 74.9 | 69.5 | 68.1 | 75 |  | 70.2385
| 1 | 5 | 73.2 | 69.2219 | 70.1359 | 68 | 74 | 71 | 70.5495 | 66.7 | 74.9 | 65.5687 | 77.2 | 65.1449 | 67.8 | 76.5 | 71.4 | 70.4 | 71.6 |  | 27
| 1 | 6 | 70.1 | 68.4012 | 66.7516 | 67 | 71 | 69 | 71.9985 | 65.8 | 74.8 | 60.1878 | 74.2 | 68.5202 | 71.8 | 72.8 | 70.6 | 68.2 | 72.5
| Mean |  | 72.2 | 68.6047 | 67.2852 | 68.5 | 73.3333 | 70.6667 | 71.3295 | 66.75 | 75.75 | 63.5199 | 73.0833 | 66.9065 | 69.2167 | 74.0833 | 71.4167 | 68.6833 | 72.4333
| Median |  | 72.5 | 68.6538 | 66.9225 | 69 | 73.5 | 71 | 71.3081 | 66.4 | 75.6 | 63.8979 | 73.1 | 66.8956 | 69.05 | 73.8 | 71.3 | 68.55 | 72.05
| Std Dev. |  | 1.7709 | 1.507 | 1.5301 | 0.83666 | 1.3663 | 1.0328 | 0.95755 | 2.2079 | 1.0368 | 5.8639 | 4.6846 | 1.3732 | 1.7882 | 1.5171 | 1.4428 | 1.1089 | 1.9765
| Rel.Std.Dev. |  | 2.453% | 2.197% | 2.274% | 1.221% | 1.863% | 1.462% | 1.342% | 3.308% | 1.369% | 9.232% | 6.41% | 2.052% | 2.583% | 2.048% | 2.02% | 1.615% | 2.729%
| PDM3 |  | 2.793% | -2.326% | -4.205% | -2.475% | 4.406% | 0.6096% | 1.553% | -4.967% | 7.847% | -9.565% | 4.05% | -4.744% | -1.455% | 5.474% | 1.677% | -2.214% | 3.125%
| Z-Score (Absolute) |  | 0.48 | 0.64 | 1.05 | 0.67 | 0.83 | 0 | 0.21 | 1.22 | 1.58 | 2.22 | 0.75 | 1.17 | 0.45 | 1.06 | 0.23 | 0.62 | 0.55
| Analytical results for Dy in OREAS 262b (Certified Value 2.37 ppm) | [19 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 2.5; 1.03: 2.5014; 1.06: 2.69; 1.07: 2.8; 1.09: 2.31; 1.10: 1.6014; 1.11: 2.21 |  | 1
| 1 | 2 | 1.01: 2.7; 1.03: 2.552; 1.06: 2.5391; 1.07: 3; 1.09: 2.23; 1.10: 1.6696; 1.11: 2.12 |  | 5
| 1 | 3 | 1.01: 2.43; 1.03: 2.6826; 1.06: 2.4491; 1.07: 3; 1.09: 2.32; 1.10: 1.6403; 1.11: 2.42 |  | 16
| 1 | 4 | 1.01: 2.57; 1.03: 2.4552; 1.06: 2.5654; 1.07: 2.8; 1.09: 2.29; 1.10: 1.5933; 1.11: 2.34 |  | 2.3735
| 1 | 5 | 1.01: 2.5; 1.03: 2.8044; 1.06: 2.5325; 1.07: 2.9; 1.09: 2.19; 1.10: 1.6001; 1.11: 2.16 |  | 28
| 1 | 6 | 1.01: 2.51; 1.03: 2.3966; 1.06: 2.5005; 1.07: 2.9; 1.09: 2.22; 1.10: 1.6598; 1.11: 2.03
| Mean |  | 1.01: 2.535; 1.03: 2.5654; 1.06: 2.5461; 1.07: 2.9; 1.09: 2.26; 1.10: 1.6274; 1.11: 2.2133
| Median |  | 1.01: 2.505; 1.03: 2.5267; 1.06: 2.5358; 1.07: 2.9; 1.09: 2.26; 1.10: 1.6208; 1.11: 2.185
| Std Dev. |  | 1.01: 0.09225; 1.03: 0.15216; 1.06: 0.081004; 1.07: 0.089443; 1.09: 0.053666; 1.10: 0.033401; 1.11: 0.14418
| Rel.Std.Dev. |  | 1.01: 3.639%; 1.03: 5.931%; 1.06: 3.182%; 1.07: 3.084%; 1.09: 2.375%; 1.10: 2.052%; 1.11: 6.514%
| PDM3 |  | 1.01: 6.806%; 1.03: 8.085%; 1.06: 7.273%; 1.07: 22.18%; 1.09: -4.78%; 1.10: -31.43%; 1.11: -6.746%
| Z-Score (Absolute) |  | 1.01: 0; 1.03: 0.07; 1.06: 0.03; 1.07: 0.89; 1.09: 0.67; 1.10: 2.23; 1.11: 0.79
| Analytical results for Er in OREAS 262b (Certified Value 1.31 ppm) | [19 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 1.48; 1.03: 1.3825; 1.06: 1.2808; 1.07: 1.6; 1.09: 1.25; 1.10: 1.4059; 1.11: 1.11 |  | 1
| 1 | 2 | 1.01: 1.39; 1.03: 1.4416; 1.06: 1.1816; 1.07: 1.7; 1.09: 1.21; 1.10: 1.4677; 1.11: 1.01 |  | 6
| 1 | 3 | 1.01: 1.46; 1.03: 1.3914; 1.06: 1.2345; 1.07: 1.8; 1.09: 1.29; 1.10: 1.4455; 1.11: 1.14 |  | 16
| 1 | 4 | 1.01: 1.43; 1.03: 1.8896; 1.06: 1.2382; 1.07: 1.6; 1.09: 1.24; 1.10: 1.3462; 1.11: 1.22 |  | 1.307
| 1 | 5 | 1.01: 1.45; 1.03: 1.4313; 1.06: 1.2981; 1.07: 1.6; 1.09: 1.19; 1.10: 1.3953; 1.11: 1.11 |  | 29
| 1 | 6 | 1.01: 1.44; 1.03: 1.3547; 1.06: 1.2767; 1.07: 1.6; 1.09: 1.22; 1.10: 1.3801; 1.11: 1.06
| Mean |  | 1.01: 1.4417; 1.03: 1.4819; 1.06: 1.2516; 1.07: 1.65; 1.09: 1.2333; 1.10: 1.4068; 1.11: 1.1083
| Median |  | 1.01: 1.445; 1.03: 1.4114; 1.06: 1.2575; 1.07: 1.6; 1.09: 1.23; 1.10: 1.4006; 1.11: 1.11
| Std Dev. |  | 1.01: 0.030605; 1.03: 0.20229; 1.06: 0.042457; 1.07: 0.083666; 1.09: 0.035024; 1.10: 0.044107; 1.11: 0.071391
| Rel.Std.Dev. |  | 1.01: 2.123%; 1.03: 13.65%; 1.06: 3.392%; 1.07: 5.071%; 1.09: 2.84%; 1.10: 3.135%; 1.11: 6.441%
| PDM3 |  | 1.01: 10.3%; 1.03: 13.38%; 1.06: -4.236%; 1.07: 26.24%; 1.09: -5.637%; 1.10: 7.633%; 1.11: -15.2%
| Z-Score (Absolute) |  | 1.01: 0.73; 1.03: 0.99; 1.06: 0.5; 1.07: NA; 1.09: 0.62; 1.10: 0.5; 1.11: 1.43
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 262b (Certified Value 0.74 ppm) | [19 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 0.74; 1.06: 0.74211; 1.07: 0.72; 1.09: 0.77; 1.10: 0.56109; 1.11: 0.72 |  | 1
| 1 | 2 | 1.01: 0.77; 1.06: 0.71246; 1.07: 0.73; 1.09: 0.73; 1.10: 0.57217; 1.11: 0.68 |  | 7
| 1 | 3 | 1.01: 0.75; 1.06: 0.72046; 1.07: 0.75; 1.09: 0.77; 1.10: 0.57111; 1.11: 0.79 |  | 16
| 1 | 4 | 1.01: 0.74; 1.06: 0.73875; 1.07: 0.73; 1.09: 0.77; 1.10: 0.5441; 1.11: 0.77 |  | 0.73702
| 1 | 5 | 1.01: 0.77; 1.06: 0.73927; 1.07: 0.74; 1.09: 0.71; 1.10: 0.56874; 1.11: 0.72 |  | 30
| 1 | 6 | 1.01: 0.78; 1.06: 0.71749; 1.07: 0.7; 1.09: 0.7; 1.10: 0.56042; 1.11: 0.69
| Mean |  | 1.01: 0.75833; 1.06: 0.72843; 1.07: 0.72833; 1.09: 0.74167; 1.10: 0.56294; 1.11: 0.72833
| Median |  | 1.01: 0.76; 1.06: 0.72961; 1.07: 0.73; 1.09: 0.75; 1.10: 0.56492; 1.11: 0.72
| Std Dev. |  | 1.01: 0.017224; 1.06: 0.013036; 1.07: 0.017224; 1.09: 0.032506; 1.10: 0.010493; 1.11: 0.043551
| Rel.Std.Dev. |  | 1.01: 2.271%; 1.06: 1.79%; 1.07: 2.365%; 1.09: 4.383%; 1.10: 1.864%; 1.11: 5.98%
| PDM3 |  | 1.01: 2.892%; 1.06: -1.166%; 1.07: -1.178%; 1.09: 0.6307%; 1.10: -23.62%; 1.11: -1.178%
| Z-Score (Absolute) |  | 1.01: 3.03; 1.03: NA; 1.06: 0; 1.07: 0; 1.09: 1.34; 1.10: 16.73; 1.11: 0
| Analytical results for Fe in OREAS 262b (Certified Value 2.46 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 2.51 | 2.5297 | 2.5261 | 2.0957 | 2.39 | 2.24 | 2.47 | 2.46 | 2.44 | 2.47 | 2.4283 | 2.54 | 2.4953 | 2.46 | 2.46 | 2.56 | 2.41 | 2.52 | 1
| 1 | 2 | 2.51 | 2.5555 | 2.4467 | 2.1001 | 2.39 | 2.57 | 2.48 | 2.49 | 2.45 | 2.54 | 2.3817 | 2.47 | 2.4712 | 2.45 | 2.46 | 2.42 | 2.36 | 2.44 | 0
| 1 | 3 | 2.45 | 2.5542 | 2.464 | 2.0501 | 2.42 | 2.68 | 2.46 | 2.49 | 2.43 | 2.53 | 2.4224 | 2.49 | 2.4213 | 2.42 | 2.43 | 2.52 | 2.42 | 2.38 | 16
| 1 | 4 | 2.52 | 2.5595 | 2.4708 | 2.0827 | 2.46 | 2.51 | 2.45 | 2.46 | 2.43 | 2.45 | 2.4014 | 2.48 | 2.4203 | 2.39 | 2.47 | 2.48 | 2.35 | 2.51 | 2.4603
| 1 | 5 | 2.52 | 2.5354 | 2.482 | 2.0882 | 2.45 | 2.57 | 2.45 | 2.4 | 2.43 | 2.41 | 2.4966 | 2.47 | 2.4807 | 2.42 | 2.49 | 2.51 | 2.37 | 2.49 | 31
| 1 | 6 | 2.44 | 2.5532 | 2.4977 | 2.0711 | 2.42 | 2.51 | 2.47 | 2.48 | 2.4 | 2.46 | 2.4746 | 2.46 | 2.4677 | 2.45 | 2.43 | 2.43 | 2.33 | 2.43
| Mean |  | 2.4917 | 2.5479 | 2.4812 | 2.0813 | 2.4217 | 2.5133 | 2.4633 | 2.4633 | 2.43 | 2.4767 | 2.4341 | 2.485 | 2.4594 | 2.4317 | 2.4567 | 2.4867 | 2.3733 | 2.4617
| Median |  | 2.51 | 2.5537 | 2.4764 | 2.0855 | 2.42 | 2.54 | 2.465 | 2.47 | 2.43 | 2.465 | 2.4253 | 2.475 | 2.4695 | 2.435 | 2.46 | 2.495 | 2.365 | 2.465
| Std Dev. |  | 0.03656 | 0.012231 | 0.027859 | 0.018368 | 0.029269 | 0.1476 | 0.012111 | 0.033862 | 0.016733 | 0.049666 | 0.043664 | 0.02881 | 0.031387 | 0.026394 | 0.023381 | 0.054283 | 0.035024 | 0.054191
| Rel.Std.Dev. |  | 1.467% | 0.4801% | 1.123% | 0.8825% | 1.209% | 5.873% | 0.4916% | 1.375% | 0.6886% | 2.005% | 1.794% | 1.159% | 1.276% | 1.085% | 0.9517% | 2.183% | 1.476% | 2.201%
| PDM3 |  | 1.276% | 3.562% | 0.8517% | -15.4% | -1.569% | 2.157% | 0.1244% | 0.1244% | -1.23% | 0.6663% | -1.062% | 1.005% | -0.03436% | -1.163% | -0.1466% | 1.073% | -3.534% | 0.05662%
| Z-Score (Absolute) |  | 0.75 | 2.19 | 0.48 | 9.79 | 1.05 | 1.31 | 0.02 | 0.02 | 0.83 | 0.36 | 0.73 | 0.58 | 0.08 | 0.79 | 0.15 | 0.62 | 2.29 | 0.02
| Analytical results for Ga in OREAS 262b (Certified Value 12.6 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 12.67 | 10.07 | 13.0015 | 13.1 | 15.6 | 14.88 | 12.8531 | 12.25 | 11.1047 | 11.7 | 13.1693 | 13.19 | 12.05 | 12.65 | 12.05 | 13.82 |  |  | 1
| 1 | 2 | 12.86 | 10.89 | 12.6521 | 12.9 | 13.1 | 14.78 | 13.4909 | 12.4 | 10.8386 | 11.9 | 13.5012 | 13.06 | 11.7 | 12.7 | 11.95 | 13.02 |  |  | 35
| 1 | 3 | 12.67 | < 10 | 12.378 | 13.6 | 13.7 | 14.75 | 13.2436 | 12.35 | 11.0152 | 11.7 | 12.0183 | 12.81 | 11.85 | 12.45 | 12 | 13.96 |  |  | 16
| 1 | 4 | 12.57 | < 10 | 12.1723 | 13.4 | 12.8 | 14.96 | 13.32 | 12.25 | 10.9518 | 12.7 | 13.2229 | 12.84 | 12 | 12.45 | 11.85 | 14.24 |  |  | 12.5909
| 1 | 5 | 12.8 | 10.14 | 12.8356 | 13.1 | 13.1 | 14.76 | 13.2988 | 12.15 | 11.327 | 12.55 | 12.2111 | 12.89 | 12.3 | 12.85 | 12.4 | 13.36 |  |  | 32
| 1 | 6 | 12.64 | 10.44 | 12.6389 | 13.3 | 13.1 | 14.7 | 13.5712 | 12.3 | 11.3497 | 12.1 | 11.6316 | 13.36 | 11.7 | 12.6 | 11.6 | 13.22
| Mean |  | 12.7017 | 10.385 | 12.6131 | 13.2333 | 13.5667 | 14.805 | 13.2963 | 12.2833 | 11.0978 | 12.1083 | 12.6257 | 13.025 | 11.9333 | 12.6167 | 11.975 | 13.6033
| Median |  | 12.67 | 10.29 | 12.6455 | 13.2 | 13.1 | 14.77 | 13.3094 | 12.275 | 11.0599 | 12 | 12.6902 | 12.975 | 11.925 | 12.625 | 11.975 | 13.59
| Std Dev. |  | 0.10759 | 0.37296 | 0.30072 | 0.25033 | 1.0386 | 0.096281 | 0.25028 | 0.08756 | 0.20559 | 0.42944 | 0.76781 | 0.21879 | 0.23166 | 0.15384 | 0.2622 | 0.47454
| Rel.Std.Dev. |  | 0.8471% | 3.591% | 2.384% | 1.892% | 7.655% | 0.6503% | 1.882% | 0.7128% | 1.852% | 3.547% | 6.081% | 1.68% | 1.941% | 1.219% | 2.19% | 3.488%
| PDM3 |  | 0.8796% | -17.52% | 0.1759% | 5.102% | 7.75% | 17.58% | 5.602% | -2.443% | -11.86% | -3.833% | 0.2765% | 3.448% | -5.223% | 0.2045% | -4.892% | 8.041%
| Z-Score (Absolute) |  | 0.09 | 4.32 | 0.01 | 0.66 | 1.01 | 2.34 | 0.72 | 0.36 | 1.63 | 0.55 | 0 | 0.43 | 0.74 | 0 | 0.69 | 1.05
| Analytical results for Gd in OREAS 262b (Certified Value 3.35 ppm) | [19 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 3.32; 1.02: 4.257; 1.03: 3.7467; 1.07: 3.6; 1.09: 3.3; 1.10: 1.5323; 1.11: 3.06 |  |  | 1
| 1 | 2 | 1.01: 3.3; 1.02: 4.3849; 1.03: 3.2912; 1.07: 3.8; 1.09: 3.18; 1.10: 1.6086; 1.11: 2.89 |  |  | 8
| 1 | 3 | 1.01: 3.29; 1.02: 4.391; 1.03: 3.3038; 1.07: 3.9; 1.09: 3.28; 1.10: 1.6371; 1.11: 3.21 |  |  | 16
| 1 | 4 | 1.01: 3.18; 1.02: 4.2538; 1.03: 3.5303; 1.07: 3.6; 1.09: 3.31; 1.10: 1.395; 1.11: 3.05 |  |  | 3.3522
| 1 | 5 | 1.01: 3.16; 1.02: 4.305; 1.03: 3.6382; 1.07: 3.7; 1.09: 3.2; 1.10: 1.5006; 1.11: 3.07 |  |  | 33
| 1 | 6 | 1.01: 3.32; 1.02: 4.1586; 1.03: 3.5163; 1.07: 3.6; 1.09: 3.25; 1.10: 1.541; 1.11: 2.97
| Mean |  | 1.01: 3.2617; 1.02: 4.2917; 1.03: 3.5044; 1.07: 3.7; 1.09: 3.2533; 1.10: 1.5358; 1.11: 3.0417
| Median |  | 1.01: 3.295; 1.02: 4.281; 1.03: 3.5233; 1.07: 3.65; 1.09: 3.265; 1.10: 1.5367; 1.11: 3.055
| Std Dev. |  | 1.01: 0.072226; 1.02: 0.088412; 1.03: 0.18061; 1.07: 0.12649; 1.09: 0.053541; 1.10: 0.085642; 1.11: 0.10741
| Rel.Std.Dev. |  | 1.01: 2.214%; 1.02: 2.06%; 1.03: 5.154%; 1.07: 3.419%; 1.09: 1.646%; 1.10: 5.576%; 1.11: 3.531%
| PDM3 |  | 1.01: -2.701%; 1.02: 28.03%; 1.03: 4.54%; 1.07: 10.37%; 1.09: -2.95%; 1.10: -54.19%; 1.11: -9.264%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 2.86; 1.03: 0.67; 1.07: 1.22; 1.09: 0.02; 1.10: 4.79; 1.11: 0.61
| Analytical results for Ge in OREAS 262b (Indicative Value 0.13 ppm) | [19 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.4 | < 0.2 | 1.5 | < 0.1 | 0.06 | 0.15 | 0.15 | 0.0422 | 0.08 | 0.19439 | 0.1 | 0.13 | 0.08 | [5 empty] | 1
| 1 | 2 | 1.3 | < 0.2 | 1.4 | < 0.1 | 0.06 | 0.15 | 0.38 | 0.0537 | 0.07 | 0.21467 | 0.1 | 0.12 | 0.07 | [5 empty] | 3
| 1 | 3 | 1 | < 0.2 | 1.4 | 0.1 | 0.07 | 0.12 | 0.18 | 0.0937 | 0.08 | 0.19417 | 0.11 | 0.13 | 0.09 | [5 empty] | 16
| 1 | 4 | 1.4 | < 0.2 | 1.6 | 0.2 | 0.07 | 0.14 | 0.08 | 0.6475 | 0.08 | 0.21252 | 0.11 | 0.13 | 0.06 | [5 empty] | 0.12766
| 1 | 5 | 1.5 | < 0.2 | 1.6 | 0.2 | 0.06 | 0.16 | 0.14 | 0.3882 | 0.08 | 0.18258 | 0.09 | 0.13 | 0.09 | [5 empty] | 9
| 1 | 6 | 1.4 | < 0.2 | 1.5 | 0.2 | 0.06 | 0.17 | 0.18 | 0.2522 | 0.08 | 0.17922 | 0.09 | 0.13 | 0.06
| Mean |  | 1.3333 |  | 1.5 | 0.175 | 0.063333 | 0.14833 | 0.185 | 0.24625 | 0.078333 | 0.19626 | 0.1 | 0.12833 | 0.075
| Median |  | 1.4 |  | 1.5 | 0.2 | 0.06 | 0.15 | 0.165 | 0.17295 | 0.08 | 0.19428 | 0.1 | 0.13 | 0.075
| Std Dev. |  | 0.17512 |  | 0.089443 | 0.05 | 0.005164 | 0.017224 | 0.10232 | 0.23798 | 0.0040825 | 0.014757 | 0.0089443 | 0.0040825 | 0.013784
| Rel.Std.Dev. |  | 13.13% |  | 5.963% | 28.57% | 8.154% | 11.61% | 55.31% | 96.64% | 5.212% | 7.519% | 8.944% | 3.181% | 18.38%
| PDM3 |  | 944.4% |  | 1075% | 37.08% | -50.39% | 16.19% | 44.92% | 92.9% | -38.64% | 53.74% | -21.67% | 0.528% | -41.25%
| Z-Score (Absolute) |  | 14.71 | 0.41 | 16.76 | 0 | 0.86 | 0.18 | 0.63 | 1.38 | 0.67 | 0.77 | 0.41 | 0.06 | 0.72
| Analytical results for Hf in OREAS 262b (Certified Value 2.54 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.77 | 3.513 | 2.597 | 2.3 | 0.3 | 2.3 | 2.7 | 2.6 | 2.0825 | 2.6 | 2.5951 | 2.68 | 2.7 | 2.8 | 2.7 | 2.3 |  |  | 1
| 1 | 2 | 2.69 | 3.3811 | 2.4636 | 2.3 | 0.1 | 2.2 | 2.7 | 2.6 | 2.2026 | 2.7 | 2.4267 | 2.68 | 2.7 | 2.9 | 2.5 | 2.14 |  |  | 21
| 1 | 3 | 2.76 | 3.4662 | 2.4311 | 2.3 | 0.2 | 2.3 | 2.8 | 2.5 | 2.2782 | 2.6 | 2.4217 | 2.57 | 2.8 | 2.9 | 2.6 | 2.31 |  |  | 16
| 1 | 4 | 2.72 | 3.2395 | 2.465 | 2.3 | 0.3 | 2.2 | 2.6 | 2.4 | 2.149 | 2.5 | 2.5089 | 2.61 | 2.9 | 2.8 | 2.8 | 2.25 |  |  | 2.5356
| 1 | 5 | 2.88 | 3.233 | 2.5519 | 2.2 | 0.5 | 2.1 | 2.6 | 2.6 | 2.0855 | 2.8 | 2.637 | 2.59 | 2.8 | 2.8 | 2.8 | 2.23 |  |  | 34
| 1 | 6 | 2.95 | 3.5115 | 2.5653 | 2.3 | 0.1 | 2.1 | 2.6 | 2.4 | 2.1949 | 2.7 | 2.6033 | 2.69 | 2.8 | 2.9 | 2.7 | 2.11
| Mean |  | 2.795 | 3.3907 | 2.5123 | 2.2833 | 0.25 | 2.2 | 2.6667 | 2.5167 | 2.1654 | 2.65 | 2.5321 | 2.6367 | 2.7833 | 2.85 | 2.6833 | 2.2233
| Median |  | 2.765 | 3.4236 | 2.5084 | 2.3 | 0.25 | 2.2 | 2.65 | 2.55 | 2.1719 | 2.65 | 2.552 | 2.645 | 2.8 | 2.85 | 2.7 | 2.24
| Std Dev. |  | 0.09975 | 0.12891 | 0.067469 | 0.040825 | 0.15166 | 0.089443 | 0.08165 | 0.098319 | 0.075503 | 0.10488 | 0.093672 | 0.052789 | 0.075277 | 0.054772 | 0.1169 | 0.082381
| Rel.Std.Dev. |  | 3.569% | 3.802% | 2.686% | 1.788% | 60.66% | 4.066% | 3.062% | 3.907% | 3.487% | 3.958% | 3.699% | 2.002% | 2.705% | 1.922% | 4.357% | 3.705%
| PDM3 |  | 10.23% | 33.72% | -0.9186% | -9.948% | -90.14% | -13.24% | 5.17% | -0.7461% | -14.6% | 4.512% | -0.137% | 3.987% | 9.771% | 12.4% | 5.827% | -12.31%
| Z-Score (Absolute) |  | 0.69 | 2.66 | 0.24 | 0.99 | 7.69 | 1.27 | 0.27 | 0.22 | 1.38 | 0.22 | 0.17 | 0.17 | 0.66 | 0.87 | 0.33 | 1.19
| Analytical results for Hg in OREAS 262b (Indicative Value 0.084 ppm) | [19 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.02: 1.10; 1.03: 1.14 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*OES/MS; 1.03: 4A*OES/MS |  |  | 3
3
| 1 | 1 | 1.01: < 1; 1.02: < 2; 1.03: 0.0825 |  |  | 1
| 1 | 2 | 1.01: < 1; 1.02: < 2; 1.03: 0.0796 |  |  | 3
| 1 | 3 | 1.01: < 1; 1.02: < 2; 1.03: 0.084 |  |  | 16
| 1 | 4 | 1.01: < 1; 1.02: < 2; 1.03: 0.0819 |  |  | 0.084283
| 1 | 5 | 1.01: < 1; 1.02: < 2; 1.03: 0.0848 |  |  | 9
| 1 | 6 | 1.01: < 1; 1.02: < 2; 1.03: 0.0929
| Mean |  | 1.03: 0.084283
| Median |  | 1.03: 0.08325
| Std Dev. |  | 1.03: 0.004591
| Rel.Std.Dev. |  | 1.03: 5.447%
| PDM3 |  | 1.03: 4.441e-14%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.81; 1.03: 0.67
| Analytical results for Ho in OREAS 262b (Certified Value 0.43 ppm) | [19 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.12: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.12: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.47; 1.02: 0.44728; 1.03: 0.4806; 1.07: 0.6; 1.09: 0.4; 1.10: 0.3668; 1.12: 0.4 |  |  | 1
| 1 | 2 | 1.01: 0.45; 1.02: 0.57186; 1.03: 0.45393; 1.07: 0.6; 1.09: 0.43; 1.10: 0.3922; 1.12: 0.38 |  |  | 10
| 1 | 3 | 1.01: 0.44; 1.02: 0.50681; 1.03: 0.4692; 1.07: 0.6; 1.09: 0.42; 1.10: 0.398; 1.12: 0.39 |  |  | 16
| 1 | 4 | 1.01: 0.47; 1.02: 0.4491; 1.03: 0.44788; 1.07: 0.6; 1.09: 0.42; 1.10: 0.3705; 1.12: 0.38 |  |  | 0.43225
| 1 | 5 | 1.01: 0.51; 1.02: 0.62612; 1.03: 0.46245; 1.07: 0.6; 1.09: 0.4; 1.10: 0.3611; 1.12: 0.4 |  |  | 35
| 1 | 6 | 1.01: 0.47; 1.02: 0.44473; 1.03: 0.46666; 1.07: 0.5; 1.09: 0.4; 1.10: 0.3779; 1.12: 0.38
| Mean |  | 1.01: 0.46833; 1.02: 0.50765; 1.03: 0.46345; 1.07: 0.58333; 1.09: 0.41167; 1.10: 0.37775; 1.12: 0.38833
| Median |  | 1.01: 0.47; 1.02: 0.47796; 1.03: 0.46456; 1.07: 0.6; 1.09: 0.41; 1.10: 0.3742; 1.12: 0.385
| Std Dev. |  | 1.01: 0.024014; 1.02: 0.076407; 1.03: 0.011589; 1.07: 0.040825; 1.09: 0.013292; 1.10: 0.014619; 1.12: 0.0098319
| Rel.Std.Dev. |  | 1.01: 5.128%; 1.02: 15.05%; 1.03: 2.501%; 1.07: 6.999%; 1.09: 3.229%; 1.10: 3.87%; 1.12: 2.532%
| PDM3 |  | 1.01: 8.348%; 1.02: 17.44%; 1.03: 7.219%; 1.07: 34.95%; 1.09: -4.762%; 1.10: -12.61%; 1.12: -10.16%
| Z-Score (Absolute) |  | 1.01: 0.52; 1.02: 1.18; 1.03: 0.44; 1.07: NA; 1.09: 0.44; 1.10: 1.01; 1.12: 0.83
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 262b (Certified Value 0.063 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.06 | 0.06 | < 0.1 | < 2 | 0.063 | 0.08 | 0.0802 | 0.07 | 0.077952 | < 0.1 | 0.071 | 0.062 | 0.067 | 0.065 |  |  |  |  | 1
| 1 | 2 | 0.06 | 0.06 | < 0.1 | < 2 | 0.068 | 0.08 | 0.0831 | 0.069 | 0.062689 | < 0.1 | 0.064 | 0.063 | 0.062 | 0.058 |  |  |  |  | 23
| 1 | 3 | 0.06 | 0.06 | < 0.1 | < 2 | 0.058 | 0.07 | 0.0808 | 0.066 | 0.066517 | < 0.1 | 0.066 | 0.063 | 0.057 | 0.061 |  |  |  |  | 16
| 1 | 4 | 0.07 | 0.06 | < 0.1 | < 2 | 0.06 | 0.07 | 0.0832 | 0.076 | 0.076235 | < 0.1 | 0.068 | 0.06 | 0.057 | 0.065 |  |  |  |  | 0.063365
| 1 | 5 | 0.06 | 0.06 | < 0.1 | < 2 | 0.06 | 0.07 | 0.0788 | 0.076 | 0.054429 | < 0.1 | 0.075 | 0.063 | 0.056 | 0.063 |  |  |  |  | 36
| 1 | 6 | 0.06 | 0.06 | < 0.1 | < 2 | 0.067 | 0.22 | 0.0861 | 0.069 | 0.052269 | < 0.1 | 0.067 | 0.062 | 0.056 | 0.058
| Mean |  | 0.061667 | 0.06 |  |  | 0.062667 | 0.098333 | 0.082033 | 0.071 | 0.065015 |  | 0.0685 | 0.062167 | 0.059167 | 0.061667
| Median |  | 0.06 | 0.06 |  |  | 0.0615 | 0.075 | 0.08195 | 0.0695 | 0.064603 |  | 0.0675 | 0.0625 | 0.057 | 0.062
| Std Dev. |  | 0.0040825 | 0 |  |  | 0.0040825 | 0.059805 | 0.002622 | 0.0040988 | 0.010728 |  | 0.003937 | 0.001169 | 0.004446 | 0.0032042
| Rel.Std.Dev. |  | 6.62% | 0% |  |  | 6.515% | 60.82% | 3.196% | 5.773% | 16.5% |  | 5.747% | 1.881% | 7.514% | 5.196%
| PDM3 |  | -2.68% | -5.31% |  |  | -1.102% | 55.19% | 29.46% | 12.05% | 2.605% |  | 8.104% | -1.891% | -6.625% | -2.68%
| Z-Score (Absolute) |  | 0.11 | 0.35 | 1.79 | 135.48 | 0.04 | 5.19 | 2.83 | 1.24 | 0.38 | 1.79 | 0.88 | 0.04 | 0.47 | 0.11
| Analytical results for K in OREAS 262b (Certified Value 1.63 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.666 | 1.7077 | 1.7153 | 1.4946 | 1.57 | 1.23 | 1.69 | 1.61 | 1.57 | 1.55 | 1.5485 | 1.69 | 1.6027 | 1.64 | 1.68 | 1.73 | 1.59 | 1.8 | 1
| 1 | 2 | 1.6614 | 1.7158 | 1.6487 | 1.4827 | 1.56 | 1.58 | 1.68 | 1.66 | 1.58 | 1.18 | 1.545 | 1.68 | 1.6178 | 1.63 | 1.77 | 1.59 | 1.56 | 1.73 | 0
| 1 | 3 | 1.6148 | 1.718 | 1.6457 | 1.452 | 1.59 | 1.83 | 1.68 | 1.66 | 1.56 | 1.58 | 1.5649 | 1.66 | 1.5924 | 1.61 | 1.65 | 1.66 | 1.6 | 1.71 | 16
| 1 | 4 | 1.672 | 1.7328 | 1.6684 | 1.4707 | 1.57 | 1.74 | 1.72 | 1.65 | 1.56 | 1.56 | 1.53 | 1.67 | 1.5934 | 1.6 | 1.76 | 1.65 | 1.54 | 1.76 | 1.6341
| 1 | 5 | 1.6727 | 1.7052 | 1.689 | 1.4824 | 1.57 | 1.79 | 1.66 | 1.63 | 1.56 | 1.55 | 1.5842 | 1.65 | 1.5795 | 1.6 | 1.76 | 1.66 | 1.58 | 1.7 | 37
| 1 | 6 | 1.6124 | 1.7234 | 1.677 | 1.463 | 1.58 | 1.78 | 1.71 | 1.67 | 1.53 | 1.43 | 1.5813 | 2.06 | 1.617 | 1.63 | 1.77 | 1.63 | 1.54 | 1.7
| Mean |  | 1.6499 | 1.7171 | 1.674 | 1.4742 | 1.5733 | 1.6583 | 1.69 | 1.6467 | 1.56 | 1.475 | 1.559 | 1.735 | 1.6005 | 1.6183 | 1.7317 | 1.6533 | 1.5683 | 1.7333
| Median |  | 1.6637 | 1.7169 | 1.6727 | 1.4766 | 1.57 | 1.76 | 1.685 | 1.655 | 1.56 | 1.55 | 1.5567 | 1.675 | 1.5981 | 1.62 | 1.76 | 1.655 | 1.57 | 1.72
| Std Dev. |  | 0.028418 | 0.010162 | 0.026123 | 0.015399 | 0.010328 | 0.2271 | 0.021909 | 0.022509 | 0.016733 | 0.15398 | 0.021515 | 0.15984 | 0.015057 | 0.017224 | 0.052694 | 0.045898 | 0.025626 | 0.039833
| Rel.Std.Dev. |  | 1.722% | 0.5918% | 1.56% | 1.045% | 0.6564% | 13.69% | 1.296% | 1.367% | 1.073% | 10.44% | 1.38% | 9.213% | 0.9408% | 1.064% | 3.043% | 2.776% | 1.634% | 2.298%
| PDM3 |  | 0.9651% | 5.08% | 2.442% | -9.783% | -3.719% | 1.482% | 3.42% | 0.7683% | -4.535% | -9.737% | -4.599% | 6.174% | -2.059% | -0.9656% | 5.97% | 1.176% | -4.025% | 6.072%
| Z-Score (Absolute) |  | 0.02 | 0.65 | 0.24 | 1.63 | 0.7 | 0.09 | 0.39 | 0.02 | 0.83 | 1.63 | 0.84 | 0.81 | 0.45 | 0.28 | 0.78 | 0.05 | 0.75 | 0.8
| Analytical results for La in OREAS 262b (Certified Value 34.6 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 34.65 | 34.6832 | 35.0309 | 31.9 | 30 | 27.8978 | 33.1 | 33.3 | 20.4112 | 33.5 | 34.9 | 35.4 | 35 | 33.5 | 26.58 |  |  |  | 1
| 1 | 2 | 34.71 | 34.6597 | 33.6487 | 37 | 28 | 31.0041 | 33.9 | 38.2 | 21.5475 | 34.6 | 34.6 | 34.4 | 35.4 | 33.2 | 26.21 |  |  |  | 24
| 1 | 3 | 33.96 | 34.0126 | 33.7404 | 38.2 | 29 | 29.994 | 33.1 | 34.4 | 21.0849 | 34.5 | 34.1 | 34.5 | 34.6 | 34.1 | 28.09 |  |  |  | 16
| 1 | 4 | 34.89 | 33.7468 | 35.3941 | 37.4 | 28 | 31.7692 | 33.2 | 33.4 | 19.4273 | 36 | 34.3 | 35.4 | 34.4 | 33.8 | 27.99 |  |  |  | 34.5743
| 1 | 5 | 35.49 | 33.9261 | 34.0535 | 37.2 | 27 | 30.0128 | 32.4 | 32.7 | 20.0582 | 36.5 | 34.7 | 36.8 | 34.8 | 36.1 | 27.31 |  |  |  | 38
| 1 | 6 | 34.62 | 34.4207 | 34.3281 | 36.2 | 27 | 30.419 | 33.6 | 32.9 | 20.4498 | 35.3 | 35 | 33.6 | 34.5 | 31.9 | 26.1
| Mean |  | 34.72 | 34.2415 | 34.3659 | 36.3167 | 28.1667 | 30.1828 | 33.2167 | 34.15 | 20.4965 | 35.0667 | 34.6 | 35.0167 | 34.7833 | 33.7667 | 27.0467
| Median |  | 34.68 | 34.2166 | 34.1908 | 37.1 | 28 | 30.2159 | 33.15 | 33.35 | 20.4305 | 34.95 | 34.65 | 34.95 | 34.7 | 33.65 | 26.945
| Std Dev. |  | 0.49283 | 0.39968 | 0.70767 | 2.2578 | 1.169 | 1.3064 | 0.51153 | 2.0695 | 0.74775 | 1.093 | 0.34641 | 1.1071 | 0.37103 | 1.3736 | 0.87883
| Rel.Std.Dev. |  | 1.419% | 1.167% | 2.059% | 6.217% | 4.15% | 4.328% | 1.54% | 6.06% | 3.648% | 3.117% | 1.001% | 3.162% | 1.067% | 4.068% | 3.249%
| PDM3 |  | 0.4214% | -0.9626% | -0.6027% | 5.039% | -18.53% | -12.7% | -3.927% | -1.227% | -40.72% | 1.424% | 0.07429% | 1.279% | 0.6046% | -2.336% | -21.77%
| Z-Score (Absolute) |  | 0.42 | 0 | 0.11 | 1.81 | 5.28 | 3.53 | 0.89 | 0.08 | 11.96 | 0.72 | 0.31 | 0.67 | 0.47 | 0.41 | 6.26
| Analytical results for Li in OREAS 262b (Certified Value 34.9 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 33.9 | 36.8021 | 35.8153 | 56.8891 | 32.7 | 35 | 40.3 | 33.7349 | 33.3 | 37.2 | 30.732 | 36.1 | 35.8182 | 34.4 | 35 | 36 | 34.8 | 38 | 1
| 1 | 2 | 33.8 | 36.1795 | 34.4743 | 56.0967 | 32.1 | 35 | 40 | 34.9657 | 33.2 | 36.4 | 30.81 | 35.2 | 35.2617 | 33.9 | 36.2 | 34.3 | 34.2 | 37 | 0
| 1 | 3 | 34.5 | 36.903 | 34.8844 | 55.1376 | 32.8 | 36 | 39.8 | 35.0352 | 32.9 | 35.3 | 31.044 | 35.3 | 35.2249 | 33.3 | 35.4 | 35.7 | 35 | 37 | 16
| 1 | 4 | 34.3 | 36.2469 | 35.3609 | 56.0297 | 32.3 | 36 | 40.9 | 34.5468 | 33 | 35.5 | 30.264 | 35.9 | 35.8594 | 33.1 | 36.3 | 35.4 | 33.9 | 39 | 34.8551
| 1 | 5 | 33.6 | 36.572 | 35.2033 | 55.9454 | 33.1 | 35 | 40.1 | 33.9223 | 32.8 | 35.4 | 31.356 | 35.4 | 34.3596 | 32.9 | 35.4 | 36.1 | 34.5 | 35 | 39
| 1 | 6 | 33.5 | 36.8364 | 35.3829 | 55.0361 | 32.3 | 36 | 39.1 | 34.9856 | 32.3 | 35.5 | 31.278 | 35.1 | 34.3669 | 35 | 35.5 | 35 | 33.6 | 34
| Mean |  | 33.9333 | 36.59 | 35.1869 | 55.8558 | 32.55 | 35.5 | 40.0333 | 34.5317 | 32.9167 | 35.8833 | 30.914 | 35.5 | 35.1485 | 33.7667 | 35.6333 | 35.4167 | 34.3333 | 36.6667
| Median |  | 33.85 | 36.6871 | 35.2821 | 55.9875 | 32.5 | 35.5 | 40.05 | 34.7563 | 32.95 | 35.5 | 30.927 | 35.35 | 35.2433 | 33.6 | 35.45 | 35.55 | 34.35 | 37
| Std Dev. |  | 0.39328 | 0.31314 | 0.46112 | 0.68582 | 0.37815 | 0.54772 | 0.59217 | 0.57521 | 0.35449 | 0.75741 | 0.40279 | 0.40497 | 0.6642 | 0.81894 | 0.50859 | 0.67946 | 0.53541 | 1.8619
| Rel.Std.Dev. |  | 1.159% | 0.8558% | 1.31% | 1.228% | 1.162% | 1.543% | 1.479% | 1.666% | 1.077% | 2.111% | 1.303% | 1.141% | 1.89% | 2.425% | 1.427% | 1.918% | 1.559% | 5.078%
| PDM3 |  | -2.645% | 4.977% | 0.9517% | 60.25% | -6.613% | 1.85% | 14.86% | -0.9278% | -5.562% | 2.95% | -11.31% | 1.85% | 0.8416% | -3.123% | 2.233% | 1.611% | -1.497% | 5.197%
| Z-Score (Absolute) |  | 0.82 | 0.77 | 0.07 | 12.28 | 1.64 | 0.12 | 2.83 | 0.46 | 1.43 | 0.35 | 2.62 | 0.12 | 0.09 | 0.92 | 0.2 | 0.07 | 0.58 | 0.82
| Analytical results for Lu in OREAS 262b (Certified Value 0.19 ppm) | [19 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.06: 1.09; 1.07: 1.10; 1.09: 1.14; 1.10: 1.22 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS | 2
2
| 1 | 1 | 1.01: 0.22; 1.02: 0.20546; 1.03: 0.20957; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0651; 1.10: 0.15 | 1
| 1 | 2 | 1.01: 0.21; 1.02: 0.21687; 1.03: 0.19357; 1.06: 0.2; 1.07: 0.18; 1.09: 0.0966; 1.10: 0.14 | 5
| 1 | 3 | 1.01: 0.2; 1.02: 0.19304; 1.03: 0.19729; 1.06: 0.3; 1.07: 0.2; 1.09: 0.0958; 1.10: 0.16 | 16
| 1 | 4 | 1.01: 0.21; 1.02: 0.21772; 1.03: 0.19922; 1.06: 0.2; 1.07: 0.2; 1.09: 0.0794; 1.10: 0.16 | 0.19321
| 1 | 5 | 1.01: 0.22; 1.02: 0.21176; 1.03: 0.20032; 1.06: 0.3; 1.07: 0.19; 1.09: 0.06; 1.10: 0.14 | 40
| 1 | 6 | 1.01: 0.21; 1.02: 0.2253; 1.03: 0.19809; 1.06: 0.2; 1.07: 0.19; 1.09: 0.1001; 1.10: 0.15
| Mean |  | 1.01: 0.21167; 1.02: 0.21169; 1.03: 0.19968; 1.06: 0.23333; 1.07: 0.195; 1.09: 0.082833; 1.10: 0.15
| Median |  | 1.01: 0.21; 1.02: 0.21431; 1.03: 0.19865; 1.06: 0.2; 1.07: 0.195; 1.09: 0.0876; 1.10: 0.15
| Std Dev. |  | 1.01: 0.0075277; 1.02: 0.011269; 1.03: 0.0053679; 1.06: 0.05164; 1.07: 0.010488; 1.09: 0.01734; 1.10: 0.0089443
| Rel.Std.Dev. |  | 1.01: 3.556%; 1.02: 5.323%; 1.03: 2.688%; 1.06: 22.13%; 1.07: 5.379%; 1.09: 20.93%; 1.10: 5.963%
| PDM3 |  | 1.01: 9.552%; 1.02: 9.565%; 1.03: 3.347%; 1.06: 20.77%; 1.07: 0.9257%; 1.09: -57.13%; 1.10: -22.36%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67; 1.03: 0.11; 1.06: NA; 1.07: 0.11; 1.09: 5.38; 1.10: 2.23
| Analytical results for Mg in OREAS 262b (Certified Value 0.557 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.5756 | 0.61742 | 0.58026 | 0.82274 | 0.546 | 0.54 | 0.58 | 0.55 | 0.52 | 0.56 | 0.55323 | 0.56 | 0.56844 | 0.56 | 0.53 | 0.59 | 0.52 | 0.569 | 1
| 1 | 2 | 0.5762 | 0.61507 | 0.56815 | 0.81076 | 0.549 | 0.55 | 0.57 | 0.57 | 0.52 | 0.57 | 0.54786 | 0.55 | 0.56785 | 0.56 | 0.55 | 0.53 | 0.51 | 0.551 | 0
| 1 | 3 | 0.5603 | 0.61091 | 0.56184 | 0.79547 | 0.552 | 0.57 | 0.56 | 0.56 | 0.52 | 0.57 | 0.55368 | 0.55 | 0.55627 | 0.55 | 0.53 | 0.55 | 0.52 | 0.54 | 16
| 1 | 4 | 0.5795 | 0.61538 | 0.56341 | 0.80884 | 0.551 | 0.55 | 0.59 | 0.56 | 0.52 | 0.56 | 0.54006 | 0.55 | 0.55519 | 0.54 | 0.55 | 0.54 | 0.51 | 0.57 | 0.55673
| 1 | 5 | 0.5787 | 0.61312 | 0.56619 | 0.81178 | 0.553 | 0.59 | 0.57 | 0.54 | 0.52 | 0.55 | 0.563 | 0.55 | 0.55953 | 0.54 | 0.54 | 0.55 | 0.51 | 0.526 | 41
| 1 | 6 | 0.5616 | 0.61987 | 0.57521 | 0.79535 | 0.554 | 0.53 | 0.56 | 0.57 | 0.51 | 0.56 | 0.55917 | 0.55 | 0.56695 | 0.55 | 0.54 | 0.53 | 0.5 | 0.546
| Mean |  | 0.57198 | 0.6153 | 0.56918 | 0.80749 | 0.55083 | 0.555 | 0.57167 | 0.55833 | 0.51833 | 0.56167 | 0.55283 | 0.55167 | 0.56237 | 0.55 | 0.54 | 0.54833 | 0.51167 | 0.55033
| Median |  | 0.5759 | 0.61523 | 0.56717 | 0.8098 | 0.5515 | 0.55 | 0.57 | 0.56 | 0.52 | 0.56 | 0.55345 | 0.55 | 0.56324 | 0.55 | 0.54 | 0.545 | 0.51 | 0.5485
| Std Dev. |  | 0.0086809 | 0.0031444 | 0.0071577 | 0.01054 | 0.0029269 | 0.021679 | 0.01169 | 0.01169 | 0.0040825 | 0.0075277 | 0.0081526 | 0.0040825 | 0.0060757 | 0.0089443 | 0.0089443 | 0.022286 | 0.0075277 | 0.017049
| Rel.Std.Dev. |  | 1.518% | 0.511% | 1.258% | 1.305% | 0.5314% | 3.906% | 2.045% | 2.094% | 0.7876% | 1.34% | 1.475% | 0.74% | 1.08% | 1.626% | 1.656% | 4.064% | 1.471% | 3.098%
| PDM3 |  | 2.74% | 10.52% | 2.236% | 45.04% | -1.059% | -0.3104% | 2.683% | 0.2883% | -6.897% | 0.8871% | -0.6999% | -0.9091% | 1.013% | -1.209% | -3.005% | -1.508% | -8.094% | -1.149%
| Z-Score (Absolute) |  | 1.5 | 5.11 | 1.27 | 21.1 | 0.26 | 0.09 | 1.48 | 0.37 | 2.96 | 0.65 | 0.09 | 0.19 | 0.7 | 0.33 | 1.16 | 0.46 | 3.52 | 0.3
| Analytical results for Mn in OREAS 262b (Certified Value 0.019 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0191 | 0.017987 | 0.019505 | < 0.1 | 0.0173 | 0.0172 | 0.0188 | 0.0198 | 0.0181 | 0.0191 | 0.017015 | 0.0195 | 0.018986 | 0.0186 | 0.0191 | 0.0196 | 0.0184 | 0.0192 | 1
| 1 | 2 | 0.0192 | 0.018548 | 0.018884 | < 0.1 | 0.0176 | 0.0189 | 0.0188 | 0.0204 | 0.018 | 0.0201 | 0.017015 | 0.019 | 0.018588 | 0.0188 | 0.0191 | 0.0189 | 0.0182 | 0.0187 | 26
| 1 | 3 | 0.0188 | 0.018023 | 0.018853 | < 0.1 | 0.0177 | 0.019 | 0.0186 | 0.0202 | 0.0181 | 0.0197 | 0.017123 | 0.0192 | 0.018429 | 0.0182 | 0.0186 | 0.0195 | 0.0186 | 0.0184 | 16
| 1 | 4 | 0.0193 | 0.018093 | 0.019036 | < 0.1 | 0.0176 | 0.0185 | 0.0185 | 0.0201 | 0.018 | 0.019 | 0.016801 | 0.0194 | 0.019085 | 0.0182 | 0.0192 | 0.0193 | 0.0179 | 0.0193 | 0.01868
| 1 | 5 | 0.0193 | 0.01797 | 0.019467 | < 0.1 | 0.0178 | 0.0187 | 0.0186 | 0.0203 | 0.0181 | 0.0186 | 0.01734 | 0.019 | 0.018234 | 0.0181 | 0.0194 | 0.0196 | 0.018 | 0.0192 | 42
| 1 | 6 | 0.0187 | 0.018087 | 0.019243 | < 0.1 | 0.0175 | 0.0185 | 0.0186 | 0.0204 | 0.0179 | 0.0193 | 0.017339 | 0.0191 | 0.018544 | 0.0191 | 0.0186 | 0.0187 | 0.0178 | 0.0183
| Mean |  | 0.019067 | 0.018118 | 0.019165 |  | 0.017583 | 0.018467 | 0.01865 | 0.0202 | 0.018033 | 0.0193 | 0.017106 | 0.0192 | 0.018644 | 0.0185 | 0.019 | 0.019267 | 0.01815 | 0.01885
| Median |  | 0.01915 | 0.018055 | 0.019139 |  | 0.0176 | 0.0186 | 0.0186 | 0.02025 | 0.01805 | 0.0192 | 0.017069 | 0.01915 | 0.018566 | 0.0184 | 0.0191 | 0.0194 | 0.0181 | 0.01895
| Std Dev. |  | 0.0002582 | 0.0002165 | 0.00028498 |  | 0.00017224 | 0.0006532 | 0.00012247 | 0.00022804 | 8.165e-05 | 0.00053292 | 0.0002094 | 0.00020976 | 0.00032865 | 0.0004 | 0.00032863 | 0.00038297 | 0.00030822 | 0.00044159
| Rel.Std.Dev. |  | 1.354% | 1.195% | 1.487% |  | 0.9796% | 3.537% | 0.6567% | 1.129% | 0.4528% | 2.761% | 1.224% | 1.093% | 1.763% | 2.162% | 1.73% | 1.988% | 1.698% | 2.343%
| PDM3 |  | 2.072% | -3.006% | 2.597% |  | -5.869% | -1.14% | -0.1583% | 8.14% | -3.46% | 3.321% | -8.426% | 2.786% | -0.1885% | -0.9613% | 1.715% | 3.143% | -2.835% | 0.9124%
| Z-Score (Absolute) |  | 0.44 | 0.88 | 0.58 | 43.6 | 1.63 | 0.4 | 0.14 | 2.02 | 1 | 0.77 | 2.29 | 0.63 | 0.15 | 0.35 | 0.35 | 0.72 | 0.84 | 0.14
| Analytical results for Mo in OREAS 262b (Certified Value 1.48 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.6 | 1.4416 | 3.5426 | 1.4 | 1 | 1.87 | < 2 | 1.42 | 1.4 | 2.4983 | 1.47 | 1.3783 | 1.25 | 1.5 | 1.58 | 1.66 | 1.9 |  | 1
| 1 | 2 | 1.6 | 1.4585 | 3.8264 | 1.3 | 2 | 1.88 | < 2 | 1.83 | 1.4 | 1.6589 | 1.32 | 1.5269 | 1.22 | 1.39 | 1.68 | 1.51 | 1.77 |  | 27
| 1 | 3 | 1.5 | 1.4815 | 3.1543 | 1.5 | 2 | 1.96 | < 2 | 1.74 | 1.3 | 1.9269 | 2.01 | 1.4509 | 1.35 | 1.37 | 1.64 | 1.58 | 1.91 |  | 16
| 1 | 4 | 1.5 | 1.3628 | 3.7582 | 1.3 | 2 | 1.93 | < 2 | 1.45 | 1.4 | 1.7088 | 1.46 | 1.3668 | 1.36 | 1.44 | 1.44 | 1.54 | 1.97 |  | 1.4827
| 1 | 5 | 1.6 | 1.3981 | 4.0386 | 1.4 | < 1 | 1.93 | < 2 | 1.42 | 1.4 | 1.7155 | 1.56 | 1.4763 | 1.44 | 1.51 | 2.33 | 1.56 | 1.92 |  | 43
| 1 | 6 | 1.5 | 1.4828 | 3.4097 | 1.4 | 2 | 2 | < 2 | 1.4 | 1.4 | 1.6452 | 1.44 | 1.4106 | 1.33 | 1.43 | 1.48 | 1.3 | 1.74
| Mean |  | 1.55 | 1.4375 | 3.6216 | 1.3833 | 1.8 | 1.9283 |  | 1.5433 | 1.3833 | 1.8589 | 1.5433 | 1.435 | 1.325 | 1.44 | 1.6917 | 1.525 | 1.8683
| Median |  | 1.55 | 1.45 | 3.6504 | 1.4 | 2 | 1.93 |  | 1.435 | 1.4 | 1.7122 | 1.465 | 1.4308 | 1.34 | 1.435 | 1.61 | 1.55 | 1.905
| Std Dev. |  | 0.054772 | 0.048105 | 0.31749 | 0.075277 | 0.44721 | 0.048751 |  | 0.19002 | 0.040825 | 0.32931 | 0.24122 | 0.061406 | 0.079687 | 0.056569 | 0.32585 | 0.12128 | 0.091524
| Rel.Std.Dev. |  | 3.534% | 3.346% | 8.767% | 5.442% | 24.85% | 2.528% |  | 12.31% | 2.951% | 17.71% | 15.63% | 4.279% | 6.014% | 3.928% | 19.26% | 7.953% | 4.899%
| PDM3 |  | 4.537% | -3.048% | 144.3% | -6.704% | 21.4% | 30.05% |  | 4.087% | -6.704% | 25.37% | 4.087% | -3.222% | -10.64% | -2.882% | 14.09% | 2.851% | 26.01%
| Z-Score (Absolute) |  | NA | 0.48 | 9.45 | NA | NA | 1.75 | 2.47 | 0 | NA | 1.43 | 0 | 0.49 | 0.99 | 0.47 | 0.67 | 0.08 | 1.48
Results from laboratories 1.01, 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 262b (Certified Value 0.19 wt.%) | [19 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.1925 | 0.18059 | 0.19994 | 1.6245 | 0.184 | 0.18 | 0.22 | 0.2 | 0.18 | 0.193 | 0.1784 | 0.19 | 0.19754 | 0.19 | 0.18 | 0.22 | 0.18 | 0.229 | 1
| 1 | 2 | 0.1907 | 0.18456 | 0.19016 | 1.6168 | 0.186 | 0.19 | 0.22 | 0.21 | 0.18 | 0.193 | 0.1713 | 0.19 | 0.16118 | 0.19 | 0.19 | 0.19 | 0.18 | 0.217 | 6
| 1 | 3 | 0.1843 | 0.18089 | 0.19128 | 1.6073 | 0.188 | 0.2 | 0.22 | 0.2 | 0.18 | 0.2 | 0.1725 | 0.19 | 0.15234 | 0.19 | 0.19 | 0.19 | 0.18 | 0.229 | 16
| 1 | 4 | 0.191 | 0.18432 | 0.19403 | 1.6021 | 0.184 | 0.19 | 0.22 | 0.2 | 0.18 | 0.198 | 0.1695 | 0.2 | 0.18243 | 0.19 | 0.19 | 0.19 | 0.18 | 0.234 | 0.18978
| 1 | 5 | 0.1918 | 0.18231 | 0.19356 | 1.6139 | 0.178 | 0.2 | 0.22 | 0.2 | 0.18 | 0.197 | 0.1803 | 0.19 | 0.1821 | 0.19 | 0.19 | 0.2 | 0.18 | 0.217 | 44
| 1 | 6 | 0.1835 | 0.18217 | 0.19255 | 1.575 | 0.185 | 0.2 | 0.21 | 0.2 | 0.18 | 0.196 | 0.175 | 0.19 | 0.18951 | 0.19 | 0.19 | 0.19 | 0.18 | 0.205
| Mean |  | 0.18897 | 0.18247 | 0.19359 | 1.6066 | 0.18417 | 0.19333 | 0.21833 | 0.20167 | 0.18 | 0.19617 | 0.1745 | 0.19167 | 0.17751 | 0.19 | 0.18833 | 0.19667 | 0.18 | 0.22183
| Median |  | 0.19085 | 0.18224 | 0.19306 | 1.6106 | 0.1845 | 0.195 | 0.22 | 0.2 | 0.18 | 0.1965 | 0.17375 | 0.19 | 0.18226 | 0.19 | 0.19 | 0.19 | 0.18 | 0.223
| Std Dev. |  | 0.0039828 | 0.0016713 | 0.0034251 | 0.017302 | 0.0033714 | 0.008165 | 0.0040825 | 0.0040825 | 3.0405e-17 | 0.0027869 | 0.0042033 | 0.0040825 | 0.017266 | 3.0405e-17 | 0.0040825 | 0.012111 | 3.0405e-17 | 0.010778
| Rel.Std.Dev. |  | 2.108% | 0.9159% | 1.769% | 1.077% | 1.831% | 4.223% | 1.87% | 2.024% | 1.689e-14% | 1.421% | 2.409% | 2.13% | 9.727% | 1.6e-14% | 2.168% | 6.158% | 1.689e-14% | 4.859%
| PDM3 |  | -0.4284% | -3.849% | 2.006% | 746.6% | -2.958% | 1.873% | 15.05% | 6.264% | -5.153% | 3.366% | -8.051% | 0.9943% | -6.463% | 0.1161% | -0.7621% | 3.629% | -5.153% | 16.89%
| Z-Score (Absolute) |  | 0.17 | 0.75 | 0.25 | 127.08 | 0.6 | 0.22 | 2.47 | 0.97 | 0.97 | 0.48 | 1.47 | 0.07 | 1.2 | 0.07 | 0.22 | 0.52 | 0.97 | 2.78
| Analytical results for Nb in OREAS 262b (Certified Value 7.7 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
2
| 1 | 1 | 9.05 | 26.6207 | 4.6279 | 6.5 | 1.1 | 6.4 | 5.3582 | 8.4 | 8.7 | 6.5489 | 8.6 | 7.0531 | 8.9 | 8.6 | 9.3 | 8 | 6.1 |  | 1
| 1 | 2 | 9.24 | 24.9227 | 4.7481 | 6.5 | 1.3 | 6.7 | < 5 | 8.3 | 8.6 | 7.1816 | 8.6 | 6.8174 | 8.9 | 9.5 | 9.4 | 8 | 6 |  | 7
| 1 | 3 | 8.64 | 24.787 | 13.4764 | 6.6 | 0.9 | 7 | < 5 | 8.3 | 8.5 | 7.0691 | 8.6 | 7.4932 | 8.8 | 9.8 | 9.1 | 8.1 | 6.4 |  | 16
| 1 | 4 | 9.29 | 26.1907 | 9.5716 | 6.9 | 0.4 | 6.1 | < 5 | 8.4 | 8.1 | 6.4405 | 8.7 | 6.6256 | 8.7 | 9.8 | 9.1 | 8.2 | 6.2 |  | 7.704
| 1 | 5 | 9.53 | 26.5088 | 9.3711 | 6.4 | 0.4 | 6.4 | 5.1057 | 8.4 | 8.5 | 6.2893 | 9.2 | 7.3671 | 8.7 | 8.6 | 9.3 | 8.5 | 6.1 |  | 45
| 1 | 6 | 8.58 | 24.8772 | 8.0414 | 6.7 | 0.9 | 6.5 | < 5 | 8 | 8.5 | 6.8583 | 8.9 | 7.1258 | 8.9 | 9.7 | 9.3 | 8 | 6.1
| Mean |  | 9.055 | 25.6512 | 8.3061 | 6.6 | 0.83333 | 6.5167 | 5.232 | 8.3 | 8.4833 | 6.7313 | 8.7667 | 7.0804 | 8.8167 | 9.3333 | 9.25 | 8.1333 | 6.15
| Median |  | 9.145 | 25.5567 | 8.7063 | 6.55 | 0.9 | 6.45 | 5.232 | 8.35 | 8.5 | 6.7036 | 8.65 | 7.0894 | 8.85 | 9.6 | 9.3 | 8.05 | 6.1
| Std Dev. |  | 0.37761 | 0.87672 | 3.3385 | 0.17889 | 0.36697 | 0.30605 | 0.17854 | 0.15492 | 0.20412 | 0.3595 | 0.24221 | 0.32602 | 0.098319 | 0.5785 | 0.12247 | 0.19664 | 0.13784
| Rel.Std.Dev. |  | 4.17% | 3.418% | 40.19% | 2.71% | 44.04% | 4.696% | 3.413% | 1.866% | 2.406% | 5.341% | 2.763% | 4.605% | 1.115% | 6.198% | 1.324% | 2.418% | 2.241%
| PDM3 |  | 17.54% | 233% | 7.815% | -14.33% | -89.18% | -15.41% | -32.09% | 7.736% | 10.12% | -12.63% | 13.79% | -8.096% | 14.44% | 21.15% | 20.07% | 5.572% | -20.17%
| Z-Score (Absolute) |  | 0.49 | 11.32 | 0 | 1.11 | 4.87 | 1.16 | 3.19 | 0 | 0.12 | 1.02 | 0.3 | 0.8 | 0.34 | 0.67 | 0.62 | 0.11 | 1.4
| Analytical results for Nd in OREAS 262b (Certified Value 24 ppm) | [19 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.09; 1.07: 1.10; 1.09: 1.14; 1.10: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 25.86; 1.03: 29.7898; 1.06: 26.8; 1.07: 23; 1.09: 15.9403; 1.10: 22.1 |  | 1
| 1 | 2 | 1.01: 26.16; 1.03: 28.1528; 1.06: 28; 1.07: 22.5; 1.09: 16.7517; 1.10: 21.2 |  | 8
| 1 | 3 | 1.01: 26.17; 1.03: 28.5969; 1.06: 29.9; 1.07: 23.1; 1.09: 16.4523; 1.10: 23.3 |  | 16
| 1 | 4 | 1.01: 26.31; 1.03: 28.6897; 1.06: 28.3; 1.07: 22.6; 1.09: 15.3952; 1.10: 23 |  | 23.9658
| 1 | 5 | 1.01: 26.21; 1.03: 28.8023; 1.06: 28; 1.07: 22.6; 1.09: 15.619; 1.10: 22.3 |  | 46
| 1 | 6 | 1.01: 26.06; 1.03: 28.8375; 1.06: 28.1; 1.07: 22.1; 1.09: 16.0303; 1.10: 22.1
| Mean |  | 1.01: 26.1283; 1.03: 28.8115; 1.06: 28.1833; 1.07: 22.65; 1.09: 16.0315; 1.10: 22.3333
| Median |  | 1.01: 26.165; 1.03: 28.746; 1.06: 28.05; 1.07: 22.6; 1.09: 15.9853; 1.10: 22.2
| Std Dev. |  | 1.01: 0.15433; 1.03: 0.539; 1.06: 0.99482; 1.07: 0.36194; 1.09: 0.5058; 1.10: 0.74476
| Rel.Std.Dev. |  | 1.01: 0.5906%; 1.03: 1.871%; 1.06: 3.53%; 1.07: 1.598%; 1.09: 3.155%; 1.10: 3.335%
| PDM3 |  | 1.01: 9.024%; 1.03: 20.22%; 1.06: 17.6%; 1.07: -5.49%; 1.09: -33.11%; 1.10: -6.812%
| Z-Score (Absolute) |  | 1.01: 0.4; 1.03: 1.02; 1.06: 0.87; 1.07: 0.4; 1.09: 1.93; 1.10: 0.47
| Analytical results for Ni in OREAS 262b (Certified Value 22 ppm) | [19 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 21.5 | 22.6461 | 20.4569 | 20 | 26 | 26.6 | 20.8367 | 22.4 | 20.8 | 18.2631 | 19.6 | 22.177 | 23.1 | 20.8 | 23.3 | 21.2 | 24 |  | 1
| 1 | 2 | 23.5 | 22.1845 | 21.2621 | 20 | 24 | 25.4 | 22.05 | 22 | 21.6 | 19.0466 | 21.3 | 22.0424 | 22.2 | 19.9 | 23.8 | 23.2 | 23 |  | 30
| 1 | 3 | 21.5 | 21.4726 | 19.7732 | 21 | 24 | 25.3 | 22.2733 | 21.8 | 22.1 | 19.0334 | 20.6 | 23.2274 | 23.4 | 20.9 | 22.4 | 21 | 25 |  | 16
| 1 | 4 | 22.2 | 21.8796 | 21.6218 | 21 | 24 | 25.2 | 21.6 | 20.2 | 21.7 | 18.0953 | 22.4 | 23.0886 | 22.4 | 21.5 | 22.6 | 21.5 | 24 |  | 22.0385
| 1 | 5 | 24 | 22.2961 | 20.7494 | 20 | 24 | 25 | 22.3133 | 21.2 | 21.5 | 19.2041 | 21.1 | 21.0019 | 22.3 | 22.8 | 23.1 | 22.6 | 25 |  | 47
| 1 | 6 | 20.9 | 21.1787 | 19.8231 | 21 | 25 | 25.5 | 22.8967 | 21 | 21.6 | 18.9174 | 21.2 | 21.4364 | 22.7 | 20.3 | 21.9 | 21.9 | 23
| Mean |  | 22.2667 | 21.9429 | 20.6144 | 20.5 | 24.5 | 25.5 | 21.995 | 21.4333 | 21.55 | 18.76 | 21.0333 | 22.1623 | 22.6833 | 21.0333 | 22.85 | 21.9 | 24
| Median |  | 21.85 | 22.0321 | 20.6031 | 20.5 | 24 | 25.35 | 22.1617 | 21.5 | 21.6 | 18.9754 | 21.15 | 22.1097 | 22.55 | 20.85 | 22.85 | 21.7 | 24
| Std Dev. |  | 1.2307 | 0.54531 | 0.74982 | 0.54772 | 0.83666 | 0.56569 | 0.70606 | 0.79415 | 0.42308 | 0.46208 | 0.91797 | 0.88096 | 0.47924 | 1.0231 | 0.68337 | 0.85323 | 0.89443
| Rel.Std.Dev. |  | 5.527% | 2.485% | 3.637% | 2.672% | 3.415% | 2.218% | 3.21% | 3.705% | 1.963% | 2.463% | 4.364% | 3.975% | 2.113% | 4.864% | 2.991% | 3.896% | 3.727%
| PDM3 |  | 1.035% | -0.4336% | -6.462% | -6.981% | 11.17% | 15.71% | -0.1974% | -2.746% | -2.217% | -14.88% | -4.561% | 0.5616% | 2.926% | -4.561% | 3.682% | -0.6285% | 8.9%
| Z-Score (Absolute) |  | 0.24 | 0 | 0.99 | 1.07 | 1.9 | 2.64 | 0.04 | 0.38 | 0.29 | 2.37 | 0.68 | 0.16 | 0.55 | 0.68 | 0.67 | 0.03 | 1.53
| Analytical results for P in OREAS 262b (Certified Value 0.024 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0238 | 0.025876 | < 0.1 | 0.025 | 0.02 | 0.0278 | 0.0229 | 0.023 | 0.024 | 0.020308 | 0.025 | 0.024793 | 0.03 | 0.024 | 0.025 | 0.023 | 0.024 |  | 1
| 1 | 2 | 0.0239 | 0.024525 | < 0.1 | 0.025 | 0.02 | 0.0271 | 0.0234 | 0.024 | 0.024 | 0.020808 | 0.024 | 0.024084 | 0.02 | 0.024 | 0.024 | 0.023 | 0.023 |  | 0
| 1 | 3 | 0.0232 | 0.024846 | < 0.1 | 0.025 | 0.02 | 0.0273 | 0.0226 | 0.023 | 0.024 | 0.020107 | 0.024 | 0.024179 | 0.03 | 0.024 | 0.024 | 0.023 | 0.024 |  | 16
| 1 | 4 | 0.0241 | 0.024986 | < 0.1 | 0.025 | 0.02 | 0.0277 | 0.0233 | 0.023 | 0.024 | 0.019709 | 0.025 | 0.023309 | 0.02 | 0.025 | 0.024 | 0.022 | 0.024 |  | 0.023557
| 1 | 5 | 0.024 | 0.025015 | < 0.1 | 0.024 | 0.02 | 0.0273 | 0.0226 | 0.023 | 0.024 | 0.020506 | 0.024 | 0.023279 | 0.02 | 0.025 | 0.024 | 0.023 | 0.023 |  | 48
| 1 | 6 | 0.0231 | 0.025162 | < 0.1 | 0.024 | 0.019 | 0.0275 | 0.0229 | 0.023 | 0.023 | 0.020606 | 0.025 | 0.024562 | 0.02 | 0.024 | 0.023 | 0.022 | 0.023
| Mean |  | 0.023683 | 0.025068 |  | 0.024667 | 0.019833 | 0.02745 | 0.02295 | 0.023167 | 0.023833 | 0.020341 | 0.0245 | 0.024034 | 0.023333 | 0.024333 | 0.024 | 0.022667 | 0.0235
| Median |  | 0.02385 | 0.025001 |  | 0.025 | 0.02 | 0.0274 | 0.0229 | 0.023 | 0.024 | 0.020407 | 0.0245 | 0.024132 | 0.02 | 0.024 | 0.024 | 0.023 | 0.0235
| Std Dev. |  | 0.00042622 | 0.00045059 |  | 0.0005164 | 0.00040825 | 0.00026646 | 0.00033912 | 0.00040825 | 0.00040825 | 0.00039251 | 0.00054772 | 0.00062838 | 0.005164 | 0.0005164 | 0.00063246 | 0.0005164 | 0.00054772
| Rel.Std.Dev. |  | 1.8% | 1.797% |  | 2.094% | 2.058% | 0.9707% | 1.478% | 1.762% | 1.713% | 1.93% | 2.236% | 2.615% | 22.13% | 2.122% | 2.635% | 2.278% | 2.331%
| PDM3 |  | 0.5362% | 6.416% |  | 4.71% | -15.81% | 16.53% | -2.577% | -1.657% | 1.173% | -13.65% | 4.003% | 2.026% | -0.9496% | 3.295% | 1.88% | -3.78% | -0.242%
| Z-Score (Absolute) |  | 0.15 | 1.25 | 26.47 | 0.84 | 4.05 | 3.66 | 0.89 | 0.67 | 0 | 3.53 | 0.67 | 0.2 | 0.51 | 0.51 | 0.17 | 1.18 | 0.34
| Analytical results for Pb in OREAS 262b (Certified Value 19.4 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 20.4 | 19.1876 | 18.8538 | 20 | 21 | 22 | 18.3526 | 18.8 | 18.9 | 18.7837 | 17.9 | 19.9045 | 19.9 | 19.3 | 19.8 | 19.2 |  |  | 1
| 1 | 2 | 20.2 | 18.8178 | 13.4431 | 21 | 20 | 22 | 18.8616 | 19 | 19.8 | 18.2915 | 18.5 | 19.9268 | 19.5 | 18.8 | 19.6 | 19.3 |  |  | 31
| 1 | 3 | 20.5 | 19.5078 | 11.3781 | 20 | 20 | 21 | 18.937 | 18.6 | 19.3 | 18.9802 | 18 | 19.1706 | 19 | 19.2 | 18.9 | 19.2 |  |  | 16
| 1 | 4 | 20.3 | 18.7773 | 16.7411 | 20 | 21 | 23 | 19.358 | 18.6 | 18.6 | 18.3057 | 19.1 | 19.2045 | 19.4 | 19.1 | 19.4 | 19.2 |  |  | 19.354
| 1 | 5 | 20.5 | 18.6215 | 12.2978 | 20 | 21 | 24 | 19.3142 | 18.3 | 18.7 | 19.1585 | 19.4 | 19.2404 | 19.6 | 19.8 | 19.4 | 19.7 |  |  | 49
| 1 | 6 | 20.3 | 19.3779 | 15.1164 | 21 | 20 | 22 | 19.227 | 18.2 | 18.6 | 22.2619 | 18.5 | 19.0292 | 19.8 | 19.1 | 19.9 | 18.6
| Mean |  | 20.3667 | 19.0483 | 14.6384 | 20.3333 | 20.5 | 22.3333 | 19.0084 | 18.5833 | 18.9833 | 19.2969 | 18.5667 | 19.4127 | 19.5333 | 19.2167 | 19.5 | 19.2
| Median |  | 20.35 | 19.0027 | 14.2797 | 20 | 20.5 | 22 | 19.082 | 18.6 | 18.8 | 18.882 | 18.5 | 19.2224 | 19.55 | 19.15 | 19.5 | 19.2
| Std Dev. |  | 0.12111 | 0.35997 | 2.8269 | 0.5164 | 0.54772 | 1.0328 | 0.37951 | 0.29944 | 0.47924 | 1.4945 | 0.59217 | 0.39622 | 0.32042 | 0.33116 | 0.35777 | 0.35214
| Rel.Std.Dev. |  | 0.5946% | 1.89% | 19.31% | 2.54% | 2.672% | 4.624% | 1.997% | 1.611% | 2.525% | 7.745% | 3.189% | 2.041% | 1.64% | 1.723% | 1.835% | 1.834%
| PDM3 |  | 5.232% | -1.58% | -24.37% | 5.06% | 5.921% | 15.39% | -1.786% | -3.982% | -1.915% | -0.2951% | -4.068% | 0.3028% | 0.9264% | -0.7098% | 0.7541% | -0.7959%
| Z-Score (Absolute) |  | 2.72 | 0.51 | 11.32 | 2.64 | 3.05 | 7.54 | 0.61 | 1.65 | 0.67 | 0.1 | 1.69 | 0.38 | 0.68 | 0.1 | 0.6 | 0.14
| Analytical results for Pr in OREAS 262b (Certified Value 6.97 ppm) | [19 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 7.38; 1.03: 7.5056; 1.06: 7.216; 1.07: 7.2; 1.09: 5.96; 1.10: 3.0821; 1.11: 6.07 |  |  | 1
| 1 | 2 | 1.01: 7.48; 1.03: 7.6313; 1.06: 6.9506; 1.07: 7.7; 1.09: 6.38; 1.10: 3.1415; 1.11: 6 |  |  | 10
| 1 | 3 | 1.01: 7.16; 1.03: 7.363; 1.06: 7.1479; 1.07: 8.1; 1.09: 6.07; 1.10: 3.1056; 1.11: 6.06 |  |  | 16
| 1 | 4 | 1.01: 7.56; 1.03: 7.5378; 1.06: 7.1015; 1.07: 7.7; 1.09: 6.1; 1.10: 2.8886; 1.11: 6.23 |  |  | 6.9687
| 1 | 5 | 1.01: 7.43; 1.03: 7.4619; 1.06: 7.2328; 1.07: 7.8; 1.09: 5.78; 1.10: 2.9437; 1.11: 6.09 |  |  | 50
| 1 | 6 | 1.01: 7.6; 1.03: 7.4295; 1.06: 7.1945; 1.07: 7.7; 1.09: 5.62; 1.10: 3.1413; 1.11: 5.93
| Mean |  | 1.01: 7.435; 1.03: 7.4882; 1.06: 7.1405; 1.07: 7.7; 1.09: 5.985; 1.10: 3.0505; 1.11: 6.0633
| Median |  | 1.01: 7.455; 1.03: 7.4838; 1.06: 7.1712; 1.07: 7.7; 1.09: 6.015; 1.10: 3.0938; 1.11: 6.065
| Std Dev. |  | 1.01: 0.15719; 1.03: 0.092777; 1.06: 0.10465; 1.07: 0.28983; 1.09: 0.26546; 1.10: 0.10787; 1.11: 0.10033
| Rel.Std.Dev. |  | 1.01: 2.114%; 1.03: 1.239%; 1.06: 1.466%; 1.07: 3.764%; 1.09: 4.435%; 1.10: 3.536%; 1.11: 1.655%
| PDM3 |  | 1.01: 6.692%; 1.03: 7.455%; 1.06: 2.466%; 1.07: 10.49%; 1.09: -14.12%; 1.10: -56.23%; 1.11: -12.99%
| Z-Score (Absolute) |  | 1.01: 0.35; 1.03: 0.42; 1.06: 0; 1.07: 0.67; 1.09: 1.39; 1.10: 4.93; 1.11: 1.3
| Analytical results for Rb in OREAS 262b (Certified Value 110 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 107.78 | 119.4584 | 108 | 118 | 119.1 | 109.5 | 101.4 | 84.0215 | 103.5 | 109.1357 | 102 | 109.5 | 115 | 107.5 | 113 |  |  |  | 1
| 1 | 2 | 110.38 | 115.5644 | 107 | 111 | 119.8 | 108.5 | 70.1 | 83.2905 | 105.5 | 111.3539 | 102 | 120 | 114.5 | 108.5 | 111 |  |  |  | 33
| 1 | 3 | 109.07 | 115.6717 | 106 | 120 | 119 | 108.5 | 94.4 | 80.5098 | 105.5 | 107.8374 | 99.2 | 116.5 | 112.5 | 106.5 | 112 |  |  |  | 16
| 1 | 4 | 110.36 | 118.0946 | 108 | 114 | 121 | 110 | 97.1 | 79.8628 | 115.5 | 109.6612 | 99.9 | 114 | 110 | 105.5 | 114 |  |  |  | 110.1437
| 1 | 5 | 109.38 | 117.6259 | 108 | 118 | 118.4 | 106.5 | 95.6 | 83.3469 | 111 | 107.3828 | 99.5 | 120 | 113.5 | 113 | 110 |  |  |  | 51
| 1 | 6 | 105.2 | 117.22 | 107 | 114 | 117.4 | 107.5 | 90 | 84.7311 | 123 | 108.3465 | 103 | 113 | 113.5 | 103 | 110
| Mean |  | 108.695 | 117.2725 | 107.3333 | 115.8333 | 119.1167 | 108.4167 | 91.4333 | 82.6271 | 110.6667 | 108.9529 | 100.9333 | 115.5 | 113.1667 | 107.3333 | 111.6667
| Median |  | 109.225 | 117.423 | 107.5 | 116 | 119.05 | 108.5 | 95 | 83.3187 | 108.25 | 108.7411 | 100.95 | 115.25 | 113.5 | 107 | 111.5
| Std Dev. |  | 1.9633 | 1.4874 | 0.8165 | 3.3714 | 1.2238 | 1.2813 | 11.0885 | 1.9725 | 7.4878 | 1.4407 | 1.5921 | 4.1473 | 1.7795 | 3.3566 | 1.633
| Rel.Std.Dev. |  | 1.806% | 1.268% | 0.7607% | 2.911% | 1.027% | 1.182% | 12.13% | 2.387% | 6.766% | 1.322% | 1.577% | 3.591% | 1.572% | 3.127% | 1.462%
| PDM3 |  | -1.315% | 6.472% | -2.552% | 5.166% | 8.147% | -1.568% | -16.99% | -24.98% | 0.4748% | -1.081% | -8.362% | 4.863% | 2.745% | -2.552% | 1.383%
| Z-Score (Absolute) |  | 0.04 | 1.33 | 0.26 | 1.1 | 1.63 | 0.09 | 2.8 | 4.21 | 0.27 | 0 | 1.28 | 1.05 | 0.67 | 0.26 | 0.43
| Analytical results for Re in OREAS 262b (Certified Value < 0.002 ppm) | [19 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.0037 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | [6 empty] | 1
| 1 | 2 | < 0.002 | < 0.05 | 0.008 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | [6 empty] | 34
| 1 | 3 | < 0.002 | < 0.05 | 0.01 | < 0.002 | 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | [6 empty] | 16
| 1 | 4 | < 0.002 | < 0.05 | 0.008 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | [6 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.05 | 0.008 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | [6 empty] | 52
| 1 | 6 | < 0.002 | < 0.05 | 0.007 | < 0.002 | < 0.002 | 0.009 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002
| Mean |  | 1.06: 0.0082; 1.09: 0.002; 1.10: 0.009; 1.11: 0.0037
| Median |  | 1.06: 0.008; 1.09: 0.002; 1.10: 0.009; 1.11: 0.0037
| Std Dev. |  | 1.06: 0.0010954
| Rel.Std.Dev. |  | 1.06: 13.36%
PDM3
| Z-Score (Absolute) |  | 0.67 | 193.53 | 47.2 | 0.67 | 0.67 | 20.23 | 6.34 | 0.67 | 193.53 | 0.67 | 0.67 | 0.67
| Analytical results for S in OREAS 262b (Certified Value 0.036 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0334 | 0.041151 | 0.4519 | 0.03 | 0.04 | 0.03 | 0.04 | 0.05 | < 0.1 | 0.032376 | 0.04 | 0.033985 | 0.04 | 0.02 | 0.04 | 0.03 | 0.036 |  | 1
| 1 | 2 | 0.0339 | 0.03817 | 0.42979 | 0.031 | 0.04 | 0.03 | 0.04 | 0.05 | < 0.1 | 0.031871 | 0.04 | 0.033477 | 0.037 | 0.02 | 0.04 | 0.03 | 0.037 |  | 19
| 1 | 3 | 0.0334 | 0.0399 | 0.44485 | 0.031 | 0.03 | 0.03 | 0.04 | 0.05 | < 0.1 | 0.030945 | 0.04 | 0.033398 | 0.037 | 0.02 | 0.04 | 0.03 | 0.036 |  | 16
| 1 | 4 | 0.0342 | 0.041494 | 0.4361 | 0.031 | 0.04 | 0.03 | 0.04 | 0.05 | < 0.1 | 0.029035 | 0.04 | 0.032151 | 0.037 | 0.02 | 0.04 | 0.03 | 0.036 |  | 0.036377
| 1 | 5 | 0.0342 | 0.041078 | 0.43928 | 0.03 | 0.04 | 0.03 | 0.04 | 0.05 | < 0.1 | 0.031639 | 0.04 | 0.032395 | 0.037 | 0.03 | 0.04 | 0.03 | 0.034 |  | 53
| 1 | 6 | 0.0331 | 0.038089 | 0.43165 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | < 0.1 | 0.03135 | 0.04 | 0.033339 | 0.04 | 0.02 | 0.04 | 0.03 | 0.035
| Mean |  | 0.0337 | 0.03998 | 0.43893 | 0.0305 | 0.036667 | 0.03 | 0.04 | 0.05 |  | 0.031203 | 0.04 | 0.033124 | 0.038 | 0.021667 | 0.04 | 0.03 | 0.035667
| Median |  | 0.03365 | 0.040489 | 0.43769 | 0.0305 | 0.04 | 0.03 | 0.04 | 0.05 |  | 0.031495 | 0.04 | 0.033368 | 0.037 | 0.02 | 0.04 | 0.03 | 0.036
| Std Dev. |  | 0.00046476 | 0.0015316 | 0.0083433 | 0.00054772 | 0.005164 | 0 | 0 | 7.6012e-18 |  | 0.0011664 | 0 | 0.00070214 | 0.0015492 | 0.0040825 | 0 | 0 | 0.0010328
| Rel.Std.Dev. |  | 1.379% | 3.831% | 1.901% | 1.796% | 14.08% | 0% | 0% | 1.52e-14% |  | 3.738% | 0% | 2.12% | 4.077% | 18.84% | 0% | 0% | 2.896%
| PDM3 |  | -7.358% | 9.906% | 1107% | -16.16% | 0.797% | -17.53% | 9.96% | 37.45% |  | -14.22% | 9.96% | -8.941% | 4.462% | -40.44% | 9.96% | -17.53% | -1.952%
| Z-Score (Absolute) |  | 0.56 | 0.63 | 76.57 | 1.17 | 0 | 1.27 | 0.63 | 2.54 | 2.54 | 1.04 | 0.63 | 0.67 | 0.25 | 2.86 | 0.63 | 1.27 | 0.19
| Analytical results for Sb in OREAS 262b (Certified Value 5.41 ppm) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 5.92 | 5.1114 | 4.0709 | 5 | 0.7 | 5.76 | < 5 | 5.65 | 5.4 | 5.4808 | 5.29 | 5.6171 | 5.35 | 5.49 | 5.57 | 5.37 | 5.78 |  | 1
| 1 | 2 | 5.94 | 4.9862 | 4.865 | 5.2 | 0.7 | 5.62 | < 5 | 5.64 | 5.3 | 5.6942 | 5.29 | 5.4811 | 5.22 | 5.38 | 5.64 | 5.2 | 5.35 |  | 35
| 1 | 3 | 5.95 | 5.0848 | 4.4508 | 5.1 | 0.5 | 5.73 | < 5 | 5.63 | 5.2 | 5.8227 | 5.36 | 5.8963 | 5.18 | 5.42 | 5.53 | 5.24 | 5.78 |  | 16
| 1 | 4 | 5.91 | 5.2533 | 4.7203 | 5.3 | 0.2 | 5.72 | < 5 | 5.85 | 5.1 | 5.6979 | 5.72 | 5.4386 | 5.2 | 5.43 | 5.51 | 5.32 | 5.64 |  | 5.4055
| 1 | 5 | 5.92 | 4.9459 | 7.5179 | 5.1 | 0.3 | 5.55 | < 5 | 5.41 | 5.1 | 5.4641 | 5.59 | 5.7076 | 5.27 | 5.63 | 5.56 | 5.45 | 5.44 |  | 54
| 1 | 6 | 5.87 | 5.199 | 8.2844 | 5 | 0.6 | 5.53 | < 5 | 5.74 | 5.2 | 5.8018 | 5.42 | 5.8989 | 5.47 | 5.38 | 5.67 | 5.04 | 5.44
| Mean |  | 5.9183 | 5.0967 | 5.6516 | 5.1167 | 0.5 | 5.6517 |  | 5.6533 | 5.2167 | 5.6602 | 5.445 | 5.6733 | 5.2817 | 5.455 | 5.58 | 5.27 | 5.5717
| Median |  | 5.92 | 5.0981 | 4.7927 | 5.1 | 0.55 | 5.67 |  | 5.645 | 5.2 | 5.696 | 5.39 | 5.6624 | 5.245 | 5.425 | 5.565 | 5.28 | 5.54
| Std Dev. |  | 0.027869 | 0.11856 | 1.78 | 0.1169 | 0.20976 | 0.098675 |  | 0.14569 | 0.1169 | 0.15471 | 0.17467 | 0.19849 | 0.11053 | 0.094816 | 0.06261 | 0.14394 | 0.18723
| Rel.Std.Dev. |  | 0.4709% | 2.326% | 31.5% | 2.285% | 41.95% | 1.746% |  | 2.577% | 2.241% | 2.733% | 3.208% | 3.499% | 2.093% | 1.738% | 1.122% | 2.731% | 3.36%
| PDM3 |  | 9.488% | -5.711% | 4.552% | -5.343% | -90.75% | 4.555% |  | 4.585% | -3.493% | 4.713% | 0.7313% | 4.954% | -2.29% | 0.9163% | 3.229% | -2.506% | 3.075%
| Z-Score (Absolute) |  | 1.58 | 1.22 | 0.67 | 1.15 | 16.85 | 0.67 | 10.05 | 0.67 | 0.81 | 0.7 | 0.03 | 0.74 | 0.59 | 0 | 0.42 | 0.63 | 0.4
| Analytical results for Sc in OREAS 262b (Certified Value 7.02 ppm) | [19 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 6.9 | 7.492 | 6 | 8 | 8 | 6.7562 | 7 | 8 | 6.5657 | 5.9 | 7.5587 | 6.8 | 6.8 | 7.1 | 6.9 | 7.6 |  |  | 1
| 1 | 2 | 7.1 | 7.1684 | 6 | 8 | 7.8 | 6.993 | 6.7 | 7 | 6.4853 | 6 | 7.2143 | 6.8 | 6.9 | 7.3 | 6.7 | 7.4 |  |  | 36
| 1 | 3 | 7.4 | 7.1635 | 6 | 8 | 7.8 | 7.0189 | 6.7 | 7 | 6.5752 | 5.9 | 7.009 | 6.6 | 7 | 7.1 | 6.8 | 7.7 |  |  | 16
| 1 | 4 | 7.3 | 7.2548 | 6 | 8 | 8 | 7.0337 | 6.8 | 7 | 6.4564 | 6.6 | 7.5327 | 6.6 | 7.1 | 7 | 6.7 | 8 |  |  | 7.0173
| 1 | 5 | 7.3 | 7.3534 | 6 | 8 | 7.6 | 6.9375 | 6.8 | 7 | 6.7056 | 6.4 | 6.9412 | 6.5 | 7.2 | 7.3 | 7.1 | 7.4 |  |  | 55
| 1 | 6 | 7.1 | 7.276 | 6 | 8 | 7.5 | 7.0078 | 6.8 | 7 | 6.6555 | 6.2 | 7.2514 | 6.8 | 6.9 | 7.1 | 6.6 | 7.3
| Mean |  | 7.1833 | 7.2847 | 6 | 8 | 7.7833 | 6.9578 | 6.8 | 7.1667 | 6.5739 | 6.1667 | 7.2512 | 6.6833 | 6.9833 | 7.15 | 6.8 | 7.5667
| Median |  | 7.2 | 7.2654 | 6 | 8 | 7.8 | 7.0004 | 6.8 | 7 | 6.5704 | 6.1 | 7.2329 | 6.7 | 6.95 | 7.1 | 6.75 | 7.5
| Std Dev. |  | 0.18348 | 0.12398 | 0 | 0 | 0.20412 | 0.1042 | 0.10954 | 0.40825 | 0.095624 | 0.28752 | 0.25687 | 0.13292 | 0.1472 | 0.12247 | 0.17889 | 0.2582
| Rel.Std.Dev. |  | 2.554% | 1.702% | 0% | 0% | 2.623% | 1.498% | 1.611% | 5.696% | 1.455% | 4.662% | 3.542% | 1.989% | 2.108% | 1.713% | 2.631% | 3.412%
| PDM3 |  | 2.366% | 3.811% | -14.5% | 14% | 10.92% | -0.847% | -3.096% | 2.129% | -6.318% | -12.12% | 3.334% | -4.759% | -0.4838% | 1.891% | -3.096% | 7.829%
| Z-Score (Absolute) |  | 0.5 | 0.76 | NA | NA | 2.01 | 0.06 | 0.46 | NA | 1.03 | 2.06 | 0.67 | 0.76 | 0 | 0.42 | 0.46 | 1.47
Results from laboratories 1.07, 1.09 and 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 262b (Indicative Value 0.73 ppm) | [19 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | < 0.5 | < 1.5 | 17.0806 | < 1 | 0.7 | < 1 | < 5 | 1 | < 1 | 0.2776 | < 1 | < 2 | 1.1 | < 1 | < 1 | 1 |  |  | 1
| 1 | 2 | 0.8 | < 1.5 | 13.5818 | < 1 | 0.9 | < 1 | < 5 | 2 | < 1 | 0.1924 | < 1 | < 2 | 0.9 | 1 | < 1 | 1 |  |  | 37
| 1 | 3 | 0.6 | < 1.5 | 14.2733 | < 1 | 1.2 | < 1 | < 5 | 2 | < 1 | 0.2002 | < 1 | < 2 | 1 | 1 | < 1 | < 1 |  |  | 16
| 1 | 4 | 0.8 | < 1.5 | 13.7234 | < 1 | 1 | < 1 | < 5 | 1 | < 1 | 0.2077 | 1 | < 2 | 1.2 | < 1 | < 1 | < 1 |  |  | 0.72622
| 1 | 5 | 0.6 | < 1.5 | 13.8773 | < 1 | 1.1 | < 1 | < 5 | 1 | < 1 | 0.1968 | < 1 | < 2 | 0.9 | < 1 | < 1 | 1 |  |  | 9
| 1 | 6 | 0.6 | < 1.5 | 14.2106 | < 1 | 0.9 | < 1 | < 5 | 1 | < 1 | 0.2101 | < 1 | < 2 | 1.1 | < 1 | < 1 | < 1
| Mean |  | 1.01: 0.68; 1.06: 14.4578; 1.09: 0.96667; 1.12: 1.3333; 1.14: 0.21413; 1.15: 1; 1.17: 1.0333; 1.19: 1; 1.21: 1
| Median |  | 1.01: 0.6; 1.06: 14.044; 1.09: 0.95; 1.12: 1; 1.14: 0.20395; 1.15: 1; 1.17: 1.05; 1.19: 1; 1.21: 1
| Std Dev. |  | 1.01: 0.10954; 1.06: 1.3129; 1.09: 0.17512; 1.12: 0.5164; 1.14: 0.031789; 1.17: 0.12111; 1.19: 0; 1.21: 0
| Rel.Std.Dev. |  | 1.01: 16.11%; 1.06: 9.081%; 1.09: 18.12%; 1.12: 38.73%; 1.14: 14.85%; 1.17: 11.72%; 1.19: 0%; 1.21: 0%
| PDM3 |  | 1.01: -6.365%; 1.06: 1891%; 1.09: 33.11%; 1.12: 83.6%; 1.14: -70.51%; 1.15: 37.7%; 1.17: 42.29%; 1.19: 37.7%; 1.21: 37.7%
| Z-Score (Absolute) |  | 0.32 | 0.13 | 44.5 | 0.67 | 0.84 | 0.67 | 5.8 | 2.02 | 0.67 | 1.6 | 0.4 | 0.94 | 1.05 | 0.13 | 0.67 | 0.13
| Analytical results for Sm in OREAS 262b (Certified Value 4.35 ppm) | [19 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 4.69; 1.03: 4.7834; 1.06: 5.0492; 1.07: 4.1; 1.09: 4.33; 1.10: 3.7876; 1.11: 4.13 |  |  | 1
| 1 | 2 | 1.01: 4.61; 1.03: 4.7548; 1.06: 4.646; 1.07: 4.5; 1.09: 4.13; 1.10: 3.9853; 1.11: 4 |  |  | 7
| 1 | 3 | 1.01: 4.7; 1.03: 4.6305; 1.06: 4.6653; 1.07: 4.1; 1.09: 4.34; 1.10: 3.8197; 1.11: 3.99 |  |  | 16
| 1 | 4 | 1.01: 4.58; 1.03: 4.5146; 1.06: 4.643; 1.07: 4.5; 1.09: 4.33; 1.10: 3.8705; 1.11: 4.12 |  |  | 4.3519
| 1 | 5 | 1.01: 4.66; 1.03: 4.6019; 1.06: 4.8789; 1.07: 4.6; 1.09: 3.89; 1.10: 3.9083; 1.11: 4.17 |  |  | 56
| 1 | 6 | 1.01: 4.64; 1.03: 4.5459; 1.06: 4.8118; 1.07: 4.1; 1.09: 3.88; 1.10: 4.2081; 1.11: 3.92
| Mean |  | 1.01: 4.6467; 1.03: 4.6385; 1.06: 4.7824; 1.07: 4.3167; 1.09: 4.15; 1.10: 3.9299; 1.11: 4.055
| Median |  | 1.01: 4.65; 1.03: 4.6162; 1.06: 4.7385; 1.07: 4.3; 1.09: 4.23; 1.10: 3.8894; 1.11: 4.06
| Std Dev. |  | 1.01: 0.046332; 1.03: 0.10939; 1.06: 0.16315; 1.07: 0.24014; 1.09: 0.21991; 1.10: 0.15288; 1.11: 0.098539
| Rel.Std.Dev. |  | 1.01: 0.9971%; 1.03: 2.358%; 1.06: 3.411%; 1.07: 5.563%; 1.09: 5.299%; 1.10: 3.89%; 1.11: 2.43%
| PDM3 |  | 1.01: 6.773%; 1.03: 6.585%; 1.06: 9.891%; 1.07: -0.8102%; 1.09: -4.64%; 1.10: -9.697%; 1.11: -6.823%
| Z-Score (Absolute) |  | 1.01: 0.69; 1.03: 0.67; 1.06: 0.98; 1.07: 0; 1.09: 0.35; 1.10: 0.81; 1.11: 0.55
| Analytical results for Sn in OREAS 262b (Certified Value 6.72 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 6.6 | 7.4565 | 37.7458 | 6.5 | 5 | 6.8 | < 10 | 6.6 | 6.8 | 6.1441 | 6.6 | 6.9985 | 6.9 | 6.8 | 7.1 | 6.4 | 7.1 |  | 1
| 1 | 2 | 6.6 | 7.0001 | 38.6246 | 6.3 | 5 | 6.7 | < 10 | 7 | 7 | 6.3168 | 6.7 | 6.9725 | 6.7 | 6.4 | 7.1 | 6.5 | 7 |  | 22
| 1 | 3 | 6.8 | 7.2107 | 40.7164 | 6.3 | 5 | 6.8 | < 10 | 7.2 | 6.8 | 6.4109 | 6.5 | 6.5183 | 6.7 | 6.5 | 7 | 6.4 | 7.1 |  | 16
| 1 | 4 | 6.6 | 7.0467 | 38.6777 | 6.4 | 4 | 6.8 | < 10 | 6.8 | 6.8 | 6.3631 | 7.3 | 6.6284 | 6.7 | 6.7 | 6.9 | 6.4 | 7.2 |  | 6.7236
| 1 | 5 | 7 | 6.908 | 41.5523 | 6.2 | 4 | 6.5 | < 10 | 6.8 | 6.8 | 6.1024 | 7 | 6.4538 | 6.7 | 6.8 | 7 | 6.9 | 6.9 |  | 57
| 1 | 6 | 6.5 | 7.2657 | 35.4173 | 6.4 | 5 | 6.4 | < 10 | 6.7 | 6.8 | 6.352 | 6.8 | 6.7128 | 6.8 | 6.6 | 7.1 | 6.3 | 6.9
| Mean |  | 6.6833 | 7.1479 | 38.789 | 6.35 | 4.6667 | 6.6667 |  | 6.85 | 6.8333 | 6.2816 | 6.8167 | 6.714 | 6.75 | 6.6333 | 7.0333 | 6.4833 | 7.0333
| Median |  | 6.6 | 7.1287 | 38.6512 | 6.35 | 5 | 6.75 |  | 6.8 | 6.8 | 6.3344 | 6.75 | 6.6706 | 6.7 | 6.65 | 7.05 | 6.4 | 7.05
| Std Dev. |  | 0.18348 | 0.20135 | 2.1839 | 0.10488 | 0.5164 | 0.17512 |  | 0.21679 | 0.08165 | 0.12694 | 0.29269 | 0.22851 | 0.083666 | 0.1633 | 0.08165 | 0.2137 | 0.12111
| Rel.Std.Dev. |  | 2.745% | 2.817% | 5.63% | 1.652% | 11.07% | 2.627% |  | 3.165% | 1.195% | 2.021% | 4.294% | 3.403% | 1.239% | 2.462% | 1.161% | 3.296% | 1.722%
| PDM3 |  | -0.5987% | 6.311% | 476.9% | -5.556% | -30.59% | -0.8465% |  | 1.88% | 1.632% | -6.574% | 1.384% | -0.1421% | 0.3929% | -1.342% | 4.607% | -3.573% | 4.607%
| Z-Score (Absolute) |  | 0.18 | 1.53 | 117.91 | 1.41 | NA | 0.24 | 6.37 | 0.43 | 0.37 | 1.66 | 0.31 | 0.07 | 0.07 | 0.36 | 1.11 | 0.91 | 1.11
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 262b (Certified Value 44.1 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 44.24 | 63.72 | 44.4436 | 204.6578 | 41.3 | 40.9 | 45.3 | 43.3665 | 40.3 | 49 | 35.8107 | 43.5 | 45.3521 | 42.7 | 43.5 | 45.5 | 43.9 | 40.3 | 1
| 1 | 2 | 44.55 | 64.4868 | 42.7455 | 203.1881 | 42.6 | 43.7 | 44.4 | 44.5745 | 40.4 | 45 | 36.0412 | 43.8 | 45.8077 | 42.7 | 43.6 | 45.3 | 44.1 | 39.2 | 23
| 1 | 3 | 44.28 | 61.7544 | 42.8082 | 200.9798 | 42 | 47.4 | 44.7 | 45.5975 | 40.3 | 49 | 36.3301 | 42.8 | 43.3767 | 41.5 | 44.1 | 44.9 | 43.2 | 39 | 16
| 1 | 4 | 44.46 | 65.61 | 44.269 | 203.027 | 42.5 | 45.3 | 45.4 | 45.4133 | 40.8 | 46 | 34.9127 | 46.5 | 44.6432 | 41.7 | 44.8 | 44.5 | 43.3 | 40.8 | 44.0777
| 1 | 5 | 44.32 | 62.9316 | 43.286 | 202.9685 | 40.1 | 43.6 | 43.5 | 44.0565 | 39.5 | 47 | 34.2757 | 46.1 | 42.0992 | 41.9 | 44.9 | 45 | 44.7 | 40.2 | 58
| 1 | 6 | 44.47 | 63.5688 | 44.4222 | 200.6709 | 41.7 | 41.9 | 43.6 | 45.1555 | 39.9 | 47 | 36.8325 | 44.7 | 42.5234 | 43.1 | 42.6 | 44.6 | 41.4 | 38.4
| Mean |  | 44.3867 | 63.6786 | 43.6624 | 202.582 | 41.7 | 43.8 | 44.4833 | 44.694 | 40.2 | 47.1667 | 35.7005 | 44.5667 | 43.9671 | 42.2667 | 43.9167 | 44.9667 | 43.4333 | 39.65
| Median |  | 44.39 | 63.6444 | 43.7775 | 202.9977 | 41.85 | 43.65 | 44.55 | 44.865 | 40.3 | 47 | 35.926 | 44.25 | 44.01 | 42.3 | 43.85 | 44.95 | 43.6 | 39.7
| Std Dev. |  | 0.12356 | 1.3159 | 0.80841 | 1.4995 | 0.92304 | 2.3358 | 0.81343 | 0.86368 | 0.44721 | 1.6021 | 0.94411 | 1.4801 | 1.5287 | 0.65013 | 0.87044 | 0.38816 | 1.1378 | 0.92033
| Rel.Std.Dev. |  | 0.2784% | 2.066% | 1.851% | 0.7402% | 2.214% | 5.333% | 1.829% | 1.932% | 1.112% | 3.397% | 2.645% | 3.321% | 3.477% | 1.538% | 1.982% | 0.8632% | 2.62% | 2.321%
| PDM3 |  | 0.701% | 44.47% | -0.9422% | 359.6% | -5.394% | -0.63% | 0.9203% | 1.398% | -8.797% | 7.008% | -19.01% | 1.109% | -0.251% | -4.109% | -0.3653% | 2.017% | -1.462% | -10.05%
| Z-Score (Absolute) |  | 0.34 | 14.98 | 0.21 | 120.4 | 1.7 | 0.11 | 0.41 | 0.57 | 2.84 | 2.45 | 6.25 | 0.47 | 0.02 | 1.27 | 0.02 | 0.78 | 0.39 | 3.26
| Analytical results for Ta in OREAS 262b (Certified Value 0.78 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.83 | 0.99017 | 0.59 | < 0.1 | 0.48 | < 5 | 0.81 | 0.8 | 0.6026 | 0.76 | 0.72097 | 0.8 | 0.8 | 0.88 | 0.75 | 0.95 |  |  | 1
| 1 | 2 | 0.83 | 1.1113 | 0.66 | < 0.1 | 0.48 | < 5 | 0.77 | 0.8 | 0.5976 | 0.72 | 0.72086 | 0.8 | 0.75 | 0.86 | 0.76 | 0.81 |  |  | 24
| 1 | 3 | 0.78 | 0.99709 | 0.62 | < 0.1 | 0.47 | < 5 | 0.8 | 0.7 | 0.552 | 0.74 | 0.75027 | 0.8 | 0.79 | 0.85 | 0.84 | 1.01 |  |  | 16
| 1 | 4 | 0.82 | 1.0058 | 0.6 | < 0.1 | 0.46 | < 5 | 0.81 | 0.7 | 0.5837 | 0.73 | 0.70375 | 0.8 | 0.81 | 0.82 | 0.79 | 0.86 |  |  | 0.78102
| 1 | 5 | 0.88 | 1.0328 | 0.62 | < 0.1 | 0.46 | < 5 | 0.8 | 0.7 | 0.5408 | 0.78 | 0.65905 | 0.8 | 0.8 | 0.91 | 0.8 | 0.81 |  |  | 59
| 1 | 6 | 0.85 | 1.0188 | 0.56 | < 0.1 | 0.47 | < 5 | 0.77 | 0.8 | 0.558 | 0.75 | 0.66393 | 0.8 | 0.83 | 0.82 | 0.77 | 0.89
| Mean |  | 0.83167 | 1.026 | 0.60833 |  | 0.47 |  | 0.79333 | 0.75 | 0.57245 | 0.74667 | 0.70314 | 0.8 | 0.79667 | 0.85667 | 0.785 | 0.88833
| Median |  | 0.83 | 1.0123 | 0.61 |  | 0.47 |  | 0.8 | 0.75 | 0.57085 | 0.745 | 0.7123 | 0.8 | 0.8 | 0.855 | 0.78 | 0.875
| Std Dev. |  | 0.033116 | 0.044489 | 0.033714 |  | 0.0089443 |  | 0.018619 | 0.054772 | 0.025678 | 0.021602 | 0.035594 | 1.2162e-16 | 0.026583 | 0.035024 | 0.032711 | 0.079603
| Rel.Std.Dev. |  | 3.982% | 4.336% | 5.542% |  | 1.903% |  | 2.347% | 7.303% | 4.486% | 2.893% | 5.062% | 1.52e-14% | 3.337% | 4.088% | 4.167% | 8.961%
| PDM3 |  | 6.485% | 31.37% | -22.11% |  | -39.82% |  | 1.577% | -3.971% | -26.7% | -4.398% | -9.972% | 2.431% | 2.004% | 9.686% | 0.51% | 13.74%
| Z-Score (Absolute) |  | 0.31 | 1.72 | 1.32 | 5.38 | 2.32 | 12.46 | 0.03 | NA | 1.58 | 0.31 | 0.63 | NA | 0.05 | 0.49 | 0.03 | 0.72
Results from laboratories 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 262b (Certified Value 0.45 ppm) | [19 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.07: 1.03; 1.09: 1.09; 1.10: 1.10; 1.11: 1.14; 1.12: 1.17; 1.13: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.07: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*OES/MS; 1.13: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.47; 1.02: 0.48891; 1.07: 0.44147; 1.09: 0.4; 1.10: 0.46; 1.11: < 0.2; 1.12: 0.5; 1.13: 0.4 |  |  | 1
| 1 | 2 | 1.01: 0.45; 1.02: 0.52645; 1.07: 0.4348; 1.09: 0.5; 1.10: 0.48; 1.11: < 0.2; 1.12: 0.5; 1.13: 0.39 |  |  | 25
| 1 | 3 | 1.01: 0.44; 1.02: 0.4847; 1.07: 0.42917; 1.09: 0.5; 1.10: 0.47; 1.11: < 0.2; 1.12: 0.5; 1.13: 0.4 |  |  | 16
| 1 | 4 | 1.01: 0.46; 1.02: 0.5107; 1.07: 0.4187; 1.09: 0.5; 1.10: 0.47; 1.11: < 0.2; 1.12: 0.5; 1.13: 0.39 |  |  | 0.44886
| 1 | 5 | 1.01: 0.46; 1.02: 0.49202; 1.07: 0.42429; 1.09: 0.5; 1.10: 0.46; 1.11: < 0.2; 1.12: 0.5; 1.13: 0.41 |  |  | 60
| 1 | 6 | 1.01: 0.44; 1.02: 0.49531; 1.07: 0.42943; 1.09: 0.5; 1.10: 0.46; 1.11: < 0.2; 1.12: 0.5; 1.13: 0.38
| Mean |  | 1.01: 0.45333; 1.02: 0.49968; 1.07: 0.42964; 1.09: 0.48333; 1.10: 0.46667; 1.12: 0.5; 1.13: 0.395
| Median |  | 1.01: 0.455; 1.02: 0.49367; 1.07: 0.4293; 1.09: 0.5; 1.10: 0.465; 1.12: 0.5; 1.13: 0.395
| Std Dev. |  | 1.01: 0.012111; 1.02: 0.015851; 1.07: 0.0079384; 1.09: 0.040825; 1.10: 0.008165; 1.12: 0; 1.13: 0.010488
| Rel.Std.Dev. |  | 1.01: 2.671%; 1.02: 3.172%; 1.07: 1.848%; 1.09: 8.447%; 1.10: 1.75%; 1.12: 0%; 1.13: 2.655%
| PDM3 |  | 1.01: 0.9956%; 1.02: 11.32%; 1.07: -4.282%; 1.09: 7.679%; 1.10: 3.966%; 1.12: 11.39%; 1.13: -12%
| Z-Score (Absolute) |  | 1.01: 0.22; 1.02: 1.1; 1.07: 0.22; 1.09: NA; 1.10: 0.47; 1.11: 6.43; 1.12: NA; 1.13: 0.87
Results from laboratories 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 262b (Certified Value 0.11 ppm) | [19 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | < 0.2 | 0.18232 | < 0.1 | < 0.1 | 0.1 | < 5 | 0.1 | < 0.5 | 0.1638 | 0.12 | 0.16723 | 0.09 | 0.11 | 0.1 | 0.05 |  |  |  | 1
| 1 | 2 | < 0.2 | 0.12638 | 0.1 | 0.1 | 0.11 | < 5 | 0.13 | < 0.5 | 0.1253 | 0.1 | 0.3092 | 0.11 | 0.1 | 0.11 | 0.09 |  |  |  | 26
| 1 | 3 | < 0.2 | < 0.1 | 0.1 | 0.1 | 0.08 | < 5 | 0.07 | < 0.5 | 0.1638 | 0.08 | 0.20833 | 0.09 | 0.1 | 0.12 | < 0.05 |  |  |  | 16
| 1 | 4 | < 0.2 | < 0.1 | < 0.1 | 0.1 | 0.11 | < 5 | 0.11 | < 0.5 | 0.1012 | 0.12 | 0.22973 | 0.1 | 0.12 | 0.1 | < 0.05 |  |  |  | 0.11035
| 1 | 5 | < 0.2 | 0.14236 | 0.1 | < 0.1 | 0.1 | < 5 | 0.09 | < 0.5 | 0.1687 | 0.12 | 0.31792 | 0.09 | 0.12 | 0.11 | 0.06 |  |  |  | 61
| 1 | 6 | < 0.2 | 0.12289 | < 0.1 | < 0.1 | 0.11 | < 5 | 0.09 | < 0.5 | 0.106 | 0.12 | 0.15089 | 0.13 | 0.1 | 0.11 | 0.07
| Mean |  |  | 0.14349 | 0.1 | 0.1 | 0.10167 |  | 0.098333 |  | 0.13813 | 0.11 | 0.23055 | 0.10167 | 0.10833 | 0.10833 | 0.0675
| Median |  |  | 0.13437 | 0.1 | 0.1 | 0.105 |  | 0.095 |  | 0.14455 | 0.12 | 0.21903 | 0.095 | 0.105 | 0.11 | 0.065
| Std Dev. |  |  | 0.027241 | 1.6997e-17 | 1.6997e-17 | 0.01169 |  | 0.020412 |  | 0.031026 | 0.016733 | 0.070243 | 0.016021 | 0.0098319 | 0.0075277 | 0.017078
| Rel.Std.Dev. |  |  | 18.98% | 1.7e-14% | 1.7e-14% | 11.5% |  | 20.76% |  | 22.46% | 15.21% | 30.47% | 15.76% | 9.076% | 6.949% | 25.3%
| PDM3 |  |  | 30.02% | -9.381% | -9.381% | -7.871% |  | -10.89% |  | 25.17% | -0.3193% | 108.9% | -7.871% | -1.83% | -1.83% | -38.83%
| Z-Score (Absolute) |  | 0.67 | 0.32 | NA | NA | 0.54 | 193.53 | 0.81 | 11.46 | 2.41 | 0.13 | 9.89 | 0.54 | 0 | 0 | 4.45
| Analytical results for Th in OREAS 262b (Certified Value 12.6 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 13.02 | 12.5216 | 12.2273 | 12.2 | 13.4 | 10.27 | 12.25 | 14.2 | 14.765 | 11.35 | 13.1027 | 11.52 | 13.1 | 13.25 | 12.35 | 12.87 |  |  | 1
| 1 | 2 | 12.91 | 12.456 | 12.5643 | 12.1 | 14 | 10.34 | 12.7 | 13.3 | 15.1804 | 11.95 | 13.2555 | 11.01 | 13.2 | 13.05 | 12.25 | 12.26 |  |  | 27
| 1 | 3 | 12.98 | 12.5092 | 12.5187 | 12.2 | 14.2 | 10.42 | 12.15 | 13.3 | 15.7172 | 11.95 | 12.7834 | 11.65 | 13.45 | 12.65 | 12.4 | 12.99 |  |  | 16
| 1 | 4 | 12.89 | 12.5314 | 12.5324 | 12.1 | 13.5 | 10.43 | 12.35 | 12.9 | 15.1584 | 12.6 | 12.8093 | 11.45 | 13.75 | 13.05 | 12.15 | 13.08 |  |  | 12.6147
| 1 | 5 | 13.11 | 12.4656 | 11.4428 | 12.2 | 13.5 | 10.24 | 12 | 13.4 | 14.8041 | 12.35 | 11.7535 | 11.98 | 13.55 | 12.9 | 12.8 | 12.74 |  |  | 62
| 1 | 6 | 13.18 | 12.4591 | 12.35 | 12.1 | 13.2 | 10.29 | 12.1 | 13 | 15.2088 | 11.7 | 11.8131 | 11.16 | 13.1 | 13.2 | 11.65 | 12.66
| Mean |  | 13.015 | 12.4905 | 12.2726 | 12.15 | 13.6333 | 10.3317 | 12.2583 | 13.35 | 15.139 | 11.9833 | 12.5863 | 11.4617 | 13.3583 | 13.0167 | 12.2667 | 12.7667
| Median |  | 13 | 12.4874 | 12.4344 | 12.15 | 13.5 | 10.315 | 12.2 | 13.3 | 15.1694 | 11.95 | 12.7963 | 11.485 | 13.325 | 13.05 | 12.3 | 12.805
| Std Dev. |  | 0.11327 | 0.034028 | 0.42657 | 0.054772 | 0.38297 | 0.079352 | 0.24782 | 0.45935 | 0.34447 | 0.44684 | 0.6472 | 0.34718 | 0.26724 | 0.21833 | 0.37506 | 0.29242
| Rel.Std.Dev. |  | 0.8703% | 0.2724% | 3.476% | 0.4508% | 2.809% | 0.768% | 2.022% | 3.441% | 2.275% | 3.729% | 5.142% | 3.029% | 2.001% | 1.677% | 3.058% | 2.29%
| PDM3 |  | 3.174% | -0.9844% | -2.712% | -3.684% | 8.075% | -18.1% | -2.825% | 5.829% | 20.01% | -5.005% | -0.2252% | -9.14% | 5.895% | 3.187% | -2.759% | 1.205%
| Z-Score (Absolute) |  | 0.67 | 0.07 | 0.38 | 0.55 | 1.55 | 3.12 | 0.4 | 1.15 | 3.67 | 0.78 | 0.07 | 1.52 | 1.16 | 0.68 | 0.38 | 0.32
| Analytical results for Ti in OREAS 262b (Certified Value 0.253 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.2722 | 0.2608 | 0.24508 | < 0.1 | 0.202 | 0.288 | 0.229 | 0.18 | 0.247 | 0.246 | 0.1221 | 0.268 | 0.26231 | 0.26 | 0.256 | 0.272 | 0.247 | 0.164 | 1
| 1 | 2 | 0.269 | 0.26367 | 0.23381 | 0.19349 | 0.227 | 0.301 | 0.227 | 0.19 | 0.244 | 0.251 | 0.12684 | 0.251 | 0.25553 | 0.26 | 0.258 | 0.25 | 0.246 | 0.161 | 28
| 1 | 3 | 0.2532 | 0.26221 | 0.238 | < 0.1 | 0.207 | 0.302 | 0.229 | 0.19 | 0.25 | 0.254 | 0.12323 | 0.258 | 0.26162 | 0.26 | 0.259 | 0.262 | 0.25 | 0.161 | 16
| 1 | 4 | 0.2744 | 0.26917 | 0.24337 | 0.18342 | 0.202 | 0.278 | 0.231 | 0.19 | 0.247 | 0.244 | 0.11599 | 0.267 | 0.25632 | 0.26 | 0.264 | 0.258 | 0.248 | 0.177 | 0.2525
| 1 | 5 | 0.2779 | 0.25937 | 0.24128 | < 0.1 | 0.2 | 0.27 | 0.225 | 0.18 | 0.253 | 0.255 | 0.11836 | 0.256 | 0.25125 | 0.26 | 0.256 | 0.262 | 0.256 | 0.172 | 63
| 1 | 6 | 0.2592 | 0.26617 | 0.24527 | 0.19937 | 0.207 | 0.289 | 0.221 | 0.19 | 0.249 | 0.254 | 0.12062 | 0.255 | 0.26237 | 0.27 | 0.258 | 0.252 | 0.242 | 0.161
| Mean |  | 0.26765 | 0.26356 | 0.24113 | 0.19209 | 0.2075 | 0.288 | 0.227 | 0.18667 | 0.24833 | 0.25067 | 0.12119 | 0.25917 | 0.25823 | 0.26167 | 0.2585 | 0.25933 | 0.24817 | 0.166
| Median |  | 0.2706 | 0.26294 | 0.24232 | 0.19349 | 0.2045 | 0.2885 | 0.228 | 0.19 | 0.248 | 0.2525 | 0.12136 | 0.257 | 0.25897 | 0.26 | 0.258 | 0.26 | 0.2475 | 0.1625
| Std Dev. |  | 0.0095221 | 0.0036131 | 0.0044983 | 0.0080643 | 0.009975 | 0.01257 | 0.0035777 | 0.005164 | 0.0030768 | 0.0046332 | 0.0038022 | 0.0068532 | 0.0045811 | 0.0040825 | 0.0029496 | 0.0079666 | 0.0046655 | 0.0068702
| Rel.Std.Dev. |  | 3.558% | 1.371% | 1.865% | 4.198% | 4.807% | 4.365% | 1.576% | 2.766% | 1.239% | 1.848% | 3.138% | 2.644% | 1.774% | 1.56% | 1.141% | 3.072% | 1.88% | 4.139%
| PDM3 |  | 6% | 4.382% | -4.502% | -23.92% | -17.82% | 14.06% | -10.1% | -26.07% | -1.651% | -0.7265% | -52.01% | 2.64% | 2.27% | 3.63% | 2.376% | 2.706% | -1.717% | -34.26%
| Z-Score (Absolute) |  | 0.93 | 0.72 | 0.43 | 6.6 | 2.16 | 1.98 | 1.16 | 3.23 | 0.06 | 0.06 | 6.6 | 0.5 | 0.45 | 0.63 | 0.46 | 0.51 | 0.07 | 4.29
| Analytical results for Tl in OREAS 262b (Certified Value 0.88 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.98 | 0.61206 | 0.84698 | 0.93 | 0.9 | 0.63 | 0.88 | 0.8 | 0.0789 | 0.82 | 0.87728 | 0.91 | 0.83 | 0.87 | 0.9 | 0.92 |  |  | 1
| 1 | 2 | 0.96 | 0.64694 | 0.8293 | 0.9 | 0.9 | 0.64 | 0.85 | 0.8 | 0.1864 | 0.83 | 0.87161 | 0.9 | 0.87 | 0.88 | 0.88 | 0.89 |  |  | 29
| 1 | 3 | 0.97 | 0.65949 | 0.82148 | 0.94 | 0.94 | 0.65 | 0.83 | 0.8 | 0.1706 | 0.82 | 0.86371 | 0.89 | 0.86 | 0.86 | 0.82 | 0.95 |  |  | 16
| 1 | 4 | 0.94 | 0.62124 | 0.83022 | 0.93 | 0.94 | 0.63 | 0.82 | 0.8 | 0.1584 | 0.93 | 0.8743 | 0.89 | 0.87 | 0.86 | 0.83 | 0.93 |  |  | 0.88233
| 1 | 5 | 1 | 0.62735 | 0.80361 | 0.9 | 0.92 | 0.6 | 0.84 | 0.8 | 0.1316 | 0.9 | 0.8281 | 0.92 | 0.87 | 0.87 | 0.94 | 0.94 |  |  | 64
| 1 | 6 | 0.94 | 0.5882 | 0.85522 | 0.93 | 0.91 | 0.63 | 0.84 | 0.8 | 0.2799 | 0.82 | 0.82818 | 0.92 | 0.85 | 0.9 | 0.79 | 0.87
| Mean |  | 0.965 | 0.62588 | 0.83113 | 0.92167 | 0.91833 | 0.63 | 0.84333 | 0.8 | 0.16763 | 0.85333 | 0.8572 | 0.905 | 0.85833 | 0.87333 | 0.86 | 0.91667
| Median |  | 0.965 | 0.6243 | 0.82976 | 0.93 | 0.915 | 0.63 | 0.84 | 0.8 | 0.1645 | 0.825 | 0.86766 | 0.905 | 0.865 | 0.87 | 0.855 | 0.925
| Std Dev. |  | 0.023452 | 0.025328 | 0.018362 | 0.017224 | 0.018348 | 0.016733 | 0.020656 | 1.2162e-16 | 0.066666 | 0.048854 | 0.022955 | 0.013784 | 0.016021 | 0.015055 | 0.056214 | 0.030768
| Rel.Std.Dev. |  | 2.43% | 4.047% | 2.209% | 1.869% | 1.998% | 2.656% | 2.449% | 1.52e-14% | 39.77% | 5.725% | 2.678% | 1.523% | 1.867% | 1.724% | 6.536% | 3.357%
| PDM3 |  | 9.369% | -29.07% | -5.803% | 4.458% | 4.08% | -28.6% | -4.42% | -9.331% | -81% | -3.287% | -2.849% | 2.569% | -2.72% | -1.02% | -2.531% | 3.891%
| Z-Score (Absolute) |  | 1.54 | 3.36 | 0.39 | 0.92 | 0.87 | 3.3 | 0.22 | NA | 9.98 | 0.07 | 0.02 | 0.67 | 0 | 0.22 | 0.02 | 0.84
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 262b (Certified Value 0.18 ppm) | [19 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.12: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.12: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.2; 1.02: 0.20064; 1.03: 0.17687; 1.07: 0.2; 1.09: 0.19; 1.10: 0.1536; 1.12: 0.17 |  |  | 1
| 1 | 2 | 1.01: 0.2; 1.02: 0.1975; 1.03: 0.16051; 1.07: 0.2; 1.09: 0.17; 1.10: 0.1592; 1.12: 0.16 |  |  | 8
| 1 | 3 | 1.01: 0.2; 1.02: 0.20785; 1.03: 0.1708; 1.07: 0.3; 1.09: 0.18; 1.10: 0.1629; 1.12: 0.17 |  |  | 16
| 1 | 4 | 1.01: 0.2; 1.02: 0.19389; 1.03: 0.16676; 1.07: 0.2; 1.09: 0.18; 1.10: 0.1532; 1.12: 0.17 |  |  | 0.17789
| 1 | 5 | 1.01: 0.2; 1.02: 0.19354; 1.03: 0.16917; 1.07: 0.2; 1.09: 0.16; 1.10: 0.1558; 1.12: 0.17 |  |  | 65
| 1 | 6 | 1.01: 0.2; 1.02: 0.19802; 1.03: 0.17483; 1.07: 0.2; 1.09: 0.16; 1.10: 0.1588; 1.12: 0.17
| Mean |  | 1.01: 0.2; 1.02: 0.19857; 1.03: 0.16983; 1.07: 0.21667; 1.09: 0.17333; 1.10: 0.15725; 1.12: 0.16833
| Median |  | 1.01: 0.2; 1.02: 0.19776; 1.03: 0.16999; 1.07: 0.2; 1.09: 0.175; 1.10: 0.1573; 1.12: 0.17
| Std Dev. |  | 1.01: 3.0405e-17; 1.02: 0.0052731; 1.03: 0.0058678; 1.07: 0.040825; 1.09: 0.012111; 1.10: 0.0037405; 1.12: 0.0040825
| Rel.Std.Dev. |  | 1.01: 1.52e-14%; 1.02: 2.656%; 1.03: 3.455%; 1.07: 18.84%; 1.09: 6.987%; 1.10: 2.379%; 1.12: 2.425%
| PDM3 |  | 1.01: 12.43%; 1.02: 11.63%; 1.03: -4.531%; 1.07: 21.8%; 1.09: -2.559%; 1.10: -11.6%; 1.12: -5.37%
| Z-Score (Absolute) |  | 1.01: 2.18; 1.02: 2.07; 1.03: 0.13; 1.07: NA; 1.09: 0.13; 1.10: 1.1; 1.12: 0.25
| Analytical results for U in OREAS 262b (Certified Value 1.95 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.98 | 1.974 | 1.7009 | 1.9 | 2.2 | 1.9 | 1.9 | 1.9 | 1.9496 | 1.7 | 2.0725 | 2.03 | 1.9 | 2.1 | 1.9 | 2.08 |  |  | 1
| 1 | 2 | 1.94 | 2.021 | 1.7119 | 1.92 | 2 | 1.8 | 2 | 2 | 1.9119 | 1.7 | 2.0591 | 1.98 | 1.8 | 2.1 | 1.9 | 2.06 |  |  | 31
| 1 | 3 | 1.97 | 2.0102 | 1.7161 | 1.92 | 2 | 1.9 | 1.9 | 1.9 | 1.9711 | 1.7 | 1.9472 | 1.94 | 1.9 | 2.1 | 1.9 | 2.15 |  |  | 16
| 1 | 4 | 1.98 | 2.0023 | 1.7207 | 1.9 | 1.9 | 1.9 | 1.9 | 1.9 | 1.9122 | 1.8 | 1.9713 | 1.97 | 1.9 | 2.1 | 1.9 | 2.15 |  |  | 1.9468
| 1 | 5 | 1.98 | 1.929 | 1.6058 | 1.92 | 2 | 1.8 | 1.9 | 1.9 | 1.9515 | 1.9 | 1.8581 | 2 | 2 | 2.1 | 2 | 2.01 |  |  | 66
| 1 | 6 | 1.96 | 1.9608 | 1.7676 | 1.96 | 1.9 | 1.8 | 1.9 | 1.9 | 1.9093 | 1.8 | 1.8802 | 2 | 1.9 | 2.1 | 1.8 | 2.12
| Mean |  | 1.9683 | 1.9829 | 1.7038 | 1.92 | 2 | 1.85 | 1.9167 | 1.9167 | 1.9343 | 1.7667 | 1.9647 | 1.9867 | 1.9 | 2.1 | 1.9 | 2.095
| Median |  | 1.975 | 1.9882 | 1.714 | 1.92 | 2 | 1.85 | 1.9 | 1.9 | 1.9309 | 1.75 | 1.9592 | 1.99 | 1.9 | 2.1 | 1.9 | 2.1
| Std Dev. |  | 0.016021 | 0.034774 | 0.053254 | 0.021909 | 0.10954 | 0.054772 | 0.040825 | 0.040825 | 0.026455 | 0.08165 | 0.088771 | 0.030768 | 0.063246 | 0 | 0.063246 | 0.055408
| Rel.Std.Dev. |  | 0.8139% | 1.754% | 3.126% | 1.141% | 5.477% | 2.961% | 2.13% | 2.13% | 1.368% | 4.622% | 4.518% | 1.549% | 3.329% | 0% | 3.329% | 2.645%
| PDM3 |  | 1.106% | 1.854% | -12.48% | -1.376% | 2.733% | -4.972% | -1.547% | -1.547% | -0.6434% | -9.252% | 0.9216% | 2.048% | -2.404% | 7.87% | -2.404% | 7.613%
| Z-Score (Absolute) |  | 0.57 | 0.78 | 3.11 | 0.1 | 1.01 | 1.07 | 0.15 | 0.15 | 0.1 | 2.23 | 0.52 | 0.83 | 0.38 | 2.4 | 0.38 | 2.33
| Analytical results for V in OREAS 262b (Certified Value 51 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 53 | 46.7166 | 53.756 | 44.4047 | 48 | 50 | 51 | 52.1667 | 52 | 55 | 47.2784 | 54 | 52.8196 | 49 | 53 | 56 | 52 | 53 | 1
| 1 | 2 | 52 | 47.138 | 51.6352 | 45.6067 | 48 | 51 | 52 | 53.6867 | 52 | 56 | 46.4733 | 52 | 54.3475 | 49 | 52 | 52 | 52 | 51 | 32
| 1 | 3 | 51 | 46.3932 | 51.9058 | 44.2911 | 48 | 52 | 50 | 53.8033 | 52 | 57 | 47.0768 | 53 | 53.7351 | 48 | 52 | 54 | 52 | 52 | 16
| 1 | 4 | 53 | 45.8542 | 52.7248 | 45.0259 | 48 | 48 | 52 | 54.1133 | 52 | 56 | 46.4802 | 55 | 54.0561 | 48 | 53 | 53 | 51 | 53 | 51.3458
| 1 | 5 | 55 | 46.0502 | 52.886 | 45.5618 | 48 | 49 | 50 | 53.4867 | 52 | 55 | 48.1799 | 53 | 53.5287 | 48 | 54 | 54 | 51 | 50 | 67
| 1 | 6 | 50 | 46.5794 | 52.8102 | 44.6158 | 48 | 48 | 49 | 54.17 | 52 | 56 | 47.4742 | 52 | 52.9488 | 50 | 53 | 52 | 50 | 50
| Mean |  | 52.3333 | 46.4553 | 52.6196 | 44.9177 | 48 | 49.6667 | 50.6667 | 53.5711 | 52 | 55.8333 | 47.1605 | 53.1667 | 53.5726 | 48.6667 | 52.8333 | 53.5 | 51.3333 | 51.5
| Median |  | 52.5 | 46.4863 | 52.7675 | 44.8209 | 48 | 49.5 | 50.5 | 53.745 | 52 | 56 | 47.1776 | 53 | 53.6319 | 48.5 | 53 | 53.5 | 51.5 | 51.5
| Std Dev. |  | 1.7512 | 0.46455 | 0.76006 | 0.57415 | 0 | 1.633 | 1.2111 | 0.73479 | 0 | 0.75277 | 0.64707 | 1.169 | 0.60308 | 0.8165 | 0.75277 | 1.5166 | 0.8165 | 1.3784
| Rel.Std.Dev. |  | 3.346% | 1% | 1.444% | 1.278% | 0% | 3.288% | 2.39% | 1.372% | 0% | 1.348% | 1.372% | 2.199% | 1.126% | 1.678% | 1.425% | 2.835% | 1.591% | 2.677%
| PDM3 |  | 1.923% | -9.525% | 2.481% | -12.52% | -6.516% | -3.27% | -1.323% | 4.334% | 1.274% | 8.74% | -8.151% | 3.546% | 4.337% | -5.218% | 2.897% | 4.195% | -0.02434% | 0.3003%
| Z-Score (Absolute) |  | 0.22 | 2 | 0.33 | 2.58 | 1.42 | 0.79 | 0.41 | 0.69 | 0.09 | 1.54 | 1.73 | 0.53 | 0.69 | 1.16 | 0.41 | 0.66 | 0.16 | 0.09
| Analytical results for W in OREAS 262b (Certified Value 3.07 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 3.4 | 3.3555 | 2.9514 | 4.7383 | 2.5 | 0.4 | 2.3 | < 10 | 3.1 | 2.8 | 2.8412 | 2.9 | 3.3369 | 3.3 | 3.6 | 3.4 | 3.1 | 2.9 | 1
| 1 | 2 | 3.3 | 3.333 | 2.8446 | 6.9127 | 2.4 | 0.4 | 2.2 | < 10 | 3.1 | 3.1 | 2.9924 | 3 | 3.3008 | 3.3 | 2.9 | 3.4 | 3.1 | 2.8 | 33
| 1 | 3 | 3.3 | 3.3075 | 2.8808 | 5.477 | 2.4 | 0.3 | 2.3 | < 10 | 3.2 | 2.9 | 2.8949 | 3 | 3.1957 | 3.4 | 3 | 3.3 | 3.1 | 2.8 | 16
| 1 | 4 | 3.4 | 3.5018 | 2.9955 | 3.5163 | 2.6 | 0.2 | 2.3 | < 10 | 3.2 | 2.6 | 2.8678 | 3.3 | 3.1311 | 3.2 | 3.3 | 3.3 | 3.2 | 2.7 | 3.0698
| 1 | 5 | 3.4 | 3.4217 | 2.7969 | 6.1866 | 2.3 | 0.1 | 2.1 | < 10 | 3.1 | 2.9 | 2.8815 | 3.3 | 2.9046 | 3.2 | 3.5 | 3.3 | 3.4 | 2.7 | 68
| 1 | 6 | 3.3 | 3.3567 | 2.9179 | 6.2566 | 2.4 | 0.3 | 2.1 | < 10 | 2.9 | 2.8 | 3.0379 | 3.3 | 2.8587 | 3.3 | 3.1 | 3.3 | 3.2 | 2.8
| Mean |  | 3.35 | 3.3794 | 2.8978 | 5.5146 | 2.4333 | 0.28333 | 2.2167 |  | 3.1 | 2.85 | 2.9193 | 3.1333 | 3.1213 | 3.2833 | 3.2333 | 3.3333 | 3.1833 | 2.7833
| Median |  | 3.35 | 3.3561 | 2.8994 | 5.8318 | 2.4 | 0.3 | 2.25 |  | 3.1 | 2.85 | 2.8882 | 3.15 | 3.1634 | 3.3 | 3.2 | 3.3 | 3.15 | 2.8
| Std Dev. |  | 0.054772 | 0.070952 | 0.072295 | 1.2292 | 0.10328 | 0.1169 | 0.098319 |  | 0.10954 | 0.16432 | 0.077699 | 0.18619 | 0.20013 | 0.075277 | 0.28048 | 0.05164 | 0.1169 | 0.075277
| Rel.Std.Dev. |  | 1.635% | 2.1% | 2.495% | 22.29% | 4.244% | 41.26% | 4.435% |  | 3.534% | 5.766% | 2.662% | 5.942% | 6.412% | 2.293% | 8.675% | 1.549% | 3.672% | 2.705%
| PDM3 |  | 9.129% | 10.09% | -5.601% | 79.64% | -20.73% | -90.77% | -27.79% |  | 0.9851% | -7.159% | -4.902% | 2.071% | 1.679% | 6.957% | 5.329% | 8.586% | 3.7% | -9.331%
| Z-Score (Absolute) |  | 0.66 | 0.75 | 0.68 | 7.12 | 2.07 | 8.48 | 2.72 | 5.59 | 0.08 | 0.83 | 0.62 | 0.02 | 0.02 | 0.47 | 0.32 | 0.61 | 0.17 | 1.03
| Analytical results for Y in OREAS 262b (Certified Value 12 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 12.86 | 12.9944 | 12.6188 | 12.2 | 14.5 | 11.5 | 10.6533 | 12.4 | 11.9 | 9.2333 | 11.6 | 12.6866 | 12.5 | 12.9 | 12.8 | 12.2 | 10.1 |  | 1
| 1 | 2 | 12.38 | 12.6859 | 11.7423 | 11.8 | 15.9 | 11.1 | 11 | 12.4 | 11.7 | 9.493 | 12.3 | 12.4208 | 12.6 | 12.3 | 13 | 11.2 | 9.6 |  | 34
| 1 | 3 | 14.22 | 12.8313 | 12.0743 | 12.4 | 15.9 | 11.7 | 11.2933 | 12 | 11 | 9.4431 | 12.2 | 12.0978 | 12.4 | 12.7 | 12.7 | 12 | 9.5 |  | 16
| 1 | 4 | 14.43 | 12.8345 | 11.7141 | 12.2 | 15.1 | 11.8 | 11.2733 | 12.6 | 10.5 | 8.8734 | 11.8 | 12.3335 | 12.4 | 13.5 | 12.8 | 11.4 | 10 |  | 12.0098
| 1 | 5 | 12.13 | 12.9576 | 12.1028 | 12 | 15.9 | 10.8 | 11.38 | 12.2 | 11 | 9.0833 | 13 | 11.677 | 12.4 | 13.2 | 12.7 | 12.6 | 10 |  | 69
| 1 | 6 | 12.17 | 13.1452 | 11.7419 | 12.1 | 15.1 | 11.3 | 11.09 | 13.6 | 10.3 | 9.323 | 12.8 | 11.7622 | 12.9 | 12.6 | 12.8 | 11.8 | 9.5
| Mean |  | 13.0317 | 12.9082 | 11.999 | 12.1167 | 15.4 | 11.3667 | 11.115 | 12.5333 | 11.0667 | 9.2415 | 12.2833 | 12.163 | 12.5333 | 12.8667 | 12.8 | 11.8667 | 9.7833
| Median |  | 12.62 | 12.896 | 11.9083 | 12.15 | 15.5 | 11.4 | 11.1817 | 12.4 | 11 | 9.2781 | 12.25 | 12.2157 | 12.45 | 12.8 | 12.8 | 11.9 | 9.8
| Std Dev. |  | 1.037 | 0.15926 | 0.35036 | 0.20412 | 0.58992 | 0.37771 | 0.26611 | 0.56095 | 0.63456 | 0.23288 | 0.54559 | 0.39263 | 0.19664 | 0.43205 | 0.10954 | 0.5164 | 0.27869
| Rel.Std.Dev. |  | 7.958% | 1.234% | 2.92% | 1.685% | 3.831% | 3.323% | 2.394% | 4.476% | 5.734% | 2.52% | 4.442% | 3.228% | 1.569% | 3.358% | 0.8558% | 4.352% | 2.849%
| PDM3 |  | 8.509% | 7.48% | -0.08957% | 0.8899% | 28.23% | -5.355% | -7.451% | 4.359% | -7.853% | -23.05% | 2.278% | 1.276% | 4.359% | 7.135% | 6.58% | -1.192% | -18.54%
| Z-Score (Absolute) |  | 0.83 | 0.71 | 0.16 | 0.04 | 3.1 | 0.76 | 1 | 0.35 | 1.05 | 2.8 | 0.12 | 0 | 0.35 | 0.67 | 0.61 | 0.28 | 2.28
| Analytical results for Yb in OREAS 262b (Certified Value 1.3 ppm) | [19 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.17; 1.12: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*OES/MS; 1.12: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 1.4; 1.03: 1.4028; 1.06: 1.3095; 1.07: 1.6; 1.09: 1.26; 1.10: 1.0321; 1.11: 1.3; 1.12: 1.1 |  | 1
| 1 | 2 | 1.01: 1.42; 1.03: 1.4756; 1.06: 1.1896; 1.07: 1.7; 1.09: 1.24; 1.10: 1.102; 1.11: 1.3; 1.12: 1 |  | 35
| 1 | 3 | 1.01: 1.4; 1.03: 1.4087; 1.06: 1.151; 1.07: 1.7; 1.09: 1.25; 1.10: 1.0522; 1.11: 1.3; 1.12: 1.1 |  | 16
| 1 | 4 | 1.01: 1.39; 1.03: 1.4588; 1.06: 1.2277; 1.07: 1.6; 1.09: 1.24; 1.10: 1.0022; 1.11: 1.3; 1.12: 1.1 |  | 1.2977
| 1 | 5 | 1.01: 1.41; 1.03: 1.4165; 1.06: 1.2584; 1.07: 1.7; 1.09: 1.23; 1.10: 1.0519; 1.11: 1.3; 1.12: 1.1 |  | 70
| 1 | 6 | 1.01: 1.43; 1.03: 1.3363; 1.06: 1.2606; 1.07: 1.6; 1.09: 1.23; 1.10: 1.0518; 1.11: 1.3; 1.12: 1.1
| Mean |  | 1.01: 1.4083; 1.03: 1.4165; 1.06: 1.2328; 1.07: 1.65; 1.09: 1.2417; 1.10: 1.0487; 1.11: 1.3; 1.12: 1.0833
| Median |  | 1.01: 1.405; 1.03: 1.4126; 1.06: 1.2431; 1.07: 1.65; 1.09: 1.24; 1.10: 1.0519; 1.11: 1.3; 1.12: 1.1
| Std Dev. |  | 1.01: 0.01472; 1.03: 0.048903; 1.06: 0.056365; 1.07: 0.054772; 1.09: 0.01169; 1.10: 0.032599; 1.11: 0; 1.12: 0.040825
| Rel.Std.Dev. |  | 1.01: 1.045%; 1.03: 3.452%; 1.06: 4.572%; 1.07: 3.32%; 1.09: 0.9415%; 1.10: 3.109%; 1.11: 0%; 1.12: 3.768%
| PDM3 |  | 1.01: 8.528%; 1.03: 9.155%; 1.06: -4.997%; 1.07: 27.15%; 1.09: -4.315%; 1.10: -19.19%; 1.11: 0.1801%; 1.12: -16.52%
| Z-Score (Absolute) |  | 1.01: 0.65; 1.03: 0.69; 1.06: 0.18; 1.07: 1.81; 1.09: 0.14; 1.10: 1.06; 1.11: 0.14; 1.12: 0.89
| Analytical results for Zn in OREAS 262b (Certified Value 47.2 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 45 | 50.549 | 50.2605 | 67.05 | 49 | 53 | 47 | 44.9791 | 46 | 46 | 45.2253 | 50 | 47.5357 | 48.2 | 48 | 53 | 47 | 47 | 1
| 1 | 2 | 45 | 51.231 | 48.5699 | 41.9634 | 49 | 53 | 46 | 44.6408 | 46 | 49 | 46.2663 | 50 | 48.0981 | 47.6 | 48 | 49 | 46 | 46 | 36
| 1 | 3 | 45 | 49.79 | 48.9039 | 41.1256 | 49 | 53 | 46 | 44.0568 | 45 | 48 | 46.6446 | 49 | 48.0716 | 46.5 | 47 | 51 | 48 | 45 | 16
| 1 | 4 | 45 | 49.9 | 49.1243 | 43.8156 | 49 | 53 | 46 | 43.8181 | 46 | 47 | 45.0734 | 50 | 47.509 | 46.3 | 49 | 50 | 46 | 47 | 47.1602
| 1 | 5 | 46 | 50.439 | 48.7342 | 41.3952 | 48 | 52 | 46 | 44.8469 | 46 | 46 | 46.5079 | 49 | 47.6734 | 46.3 | 49 | 51 | 47 | 44 | 71
| 1 | 6 | 44 | 52.474 | 49.4112 | 41.764 | 49 | 53 | 47 | 45.037 | 46 | 48 | 47.1622 | 49 | 47.5277 | 48.3 | 48 | 49 | 45 | 48
| Mean |  | 45 | 50.7305 | 49.1673 | 46.1856 | 48.8333 | 52.8333 | 46.3333 | 44.5631 | 45.8333 | 47.3333 | 46.1466 | 49.5 | 47.7359 | 47.2 | 48.1667 | 50.5 | 46.5 | 46.1667
| Median |  | 45 | 50.494 | 49.0141 | 41.8637 | 49 | 53 | 46 | 44.7439 | 46 | 47.5 | 46.3871 | 49.5 | 47.6046 | 47.05 | 48 | 50.5 | 46.5 | 46.5
| Std Dev. |  | 0.63246 | 0.99825 | 0.61159 | 10.2652 | 0.40825 | 0.40825 | 0.5164 | 0.50904 | 0.40825 | 1.2111 | 0.82762 | 0.54772 | 0.27663 | 0.94657 | 0.75277 | 1.5166 | 1.0488 | 1.472
| Rel.Std.Dev. |  | 1.405% | 1.968% | 1.244% | 22.23% | 0.836% | 0.7727% | 1.115% | 1.142% | 0.8907% | 2.559% | 1.793% | 1.107% | 0.5795% | 2.005% | 1.563% | 3.003% | 2.256% | 3.188%
| PDM3 |  | -4.581% | 7.571% | 4.256% | -2.066% | 3.548% | 12.03% | -1.753% | -5.507% | -2.813% | 0.3672% | -2.149% | 4.961% | 1.221% | 0.08445% | 2.134% | 7.082% | -1.4% | -2.107%
| Z-Score (Absolute) |  | 1.2 | 1.83 | 1 | 0.57 | 0.83 | 2.94 | 0.49 | 1.43 | 0.76 | 0.04 | 0.59 | 1.18 | 0.25 | 0.04 | 0.48 | 1.71 | 0.4 | 0.58
| Analytical results for Zr in OREAS 262b (Certified Value 89 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 95 | 90.5472 | 92.9747 | 310.4236 | 83.7 | 23 | 85.2 | 72.1 | 92 | 88.5 | 60.1092 | 98 | 85.1686 | 95.6 | 96.8 | 91.1 | 90.3 | 78.8 | 1
| 1 | 2 | 99.8 | 92.61 | 90.2682 | 96.1737 | 84.3 | 14 | 82.2 | 73.6567 | 87.6 | 95.3 | 60.583 | 94 | 84.4395 | 95.4 | 94.4 | 95.4 | 89.8 | 77.2 | 37
| 1 | 3 | 91.9 | 89.694 | 91.8719 | 92.1134 | 84.1 | 19 | 83.3 | 73.4133 | 88.9 | 90.1 | 60.9376 | 93.8 | 85.7642 | 96.2 | 97.1 | 95.1 | 89.9 | 80.3 | 16
| 1 | 4 | 93.5 | 89.37 | 93.5148 | 90.8672 | 84.7 | 25 | 82.6 | 74.4467 | 91 | 88 | 60.3696 | 99.5 | 84.8752 | 94.4 | 102.5 | 91.4 | 91.3 | 82.2 | 89.1404
| 1 | 5 | 95.5 | 92.6856 | 89.007 | 88.4663 | 87.7 | 33 | 80.4 | 74.9767 | 91.1 | 90.9 | 60.2769 | 101 | 84.2714 | 94.9 | 94.4 | 93.1 | 94.8 | 75.3 | 72
| 1 | 6 | 91.1 | 92.4264 | 91.5477 | 87.0286 | 82.2 | 12 | 86.8 | 72.2667 | 88.7 | 86.8 | 60.3018 | 98 | 85.698 | 98.8 | 101 | 94.5 | 87.1 | 75.4
| Mean |  | 94.4667 | 91.2222 | 91.5307 | 127.5121 | 84.45 | 21 | 83.4167 | 73.4767 | 89.8833 | 89.9333 | 60.4297 | 97.3833 | 85.0362 | 95.8833 | 97.7 | 93.4333 | 90.5333 | 78.2
| Median |  | 94.25 | 91.4868 | 91.7098 | 91.4903 | 84.2 | 21 | 82.95 | 73.535 | 89.95 | 89.3 | 60.3357 | 98 | 85.0219 | 95.5 | 96.95 | 93.8 | 90.1 | 78
| Std Dev. |  | 3.1194 | 1.5323 | 1.6794 | 89.6639 | 1.8108 | 7.7201 | 2.2754 | 1.1477 | 1.7198 | 3.0124 | 0.29243 | 2.9192 | 0.62489 | 1.5549 | 3.3728 | 1.8694 | 2.5129 | 2.7575
| Rel.Std.Dev. |  | 3.302% | 1.68% | 1.835% | 70.32% | 2.144% | 36.76% | 2.728% | 1.562% | 1.913% | 3.35% | 0.4839% | 2.998% | 0.7349% | 1.622% | 3.452% | 2.001% | 2.776% | 3.526%
| PDM3 |  | 5.975% | 2.335% | 2.682% | 43.05% | -5.262% | -76.44% | -6.421% | -17.57% | 0.8335% | 0.8896% | -32.21% | 9.247% | -4.604% | 7.564% | 9.602% | 4.816% | 1.563% | -12.27%
| Z-Score (Absolute) |  | 0.5 | 0.12 | 0.15 | 4.4 | 0.68 | 8.17 | 0.8 | 1.98 | 0.04 | 0.04 | 3.52 | 0.84 | 0.61 | 0.67 | 0.88 | 0.38 | 0.04 | 1.42
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C965: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1020: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1038: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 262b (Certified Value 0.132 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.13 | < 0.2 | < 10 | 0.13 | 0.14 | 0.12 | < 0.5 | 0.15 | 0.2 | < 0.3 | 0.13 | 0.15442 | 0.1 | 0.13 | 0.14 | 0.14 |  |  |  | 1
| 1 | 2 | 0.13 | < 0.2 | < 10 | 0.12 | 0.123 | 0.12 | < 0.5 | 0.14 | 0.2 | < 0.3 | 0.13 | 0.19493 | 0.1 | 0.13 | 0.13 | 0.13 |  |  |  | 21
| 1 | 3 | 0.13 | < 0.2 | < 10 | 0.13 | 0.125 | 0.1 | < 0.5 | 0.13 | 0.1 | < 0.3 | 0.11 | 0.23738 | 0.1 | 0.14 | 0.14 | 0.15 |  |  |  | 16
| 1 | 4 | 0.13 | < 0.2 | < 10 | 0.13 | 0.111 | 0.11 | < 0.5 | 0.14 | 0.1 | < 0.3 | 0.13 | 0.13698 | 0.1 | 0.13 | 0.13 | 0.14 |  |  |  | 0.13212
| 1 | 5 | 0.14 | < 0.2 | < 10 | 0.12 | 0.146 | 0.12 | < 0.5 | 0.14 | 0.2 | < 0.3 | 0.13 | 0.16031 | 0.1 | 0.14 | 0.13 | 0.13 |  |  |  | 74
| 1 | 6 | 0.13 | < 0.2 | < 10 | 0.13 | 0.137 | 0.11 | < 0.5 | 0.13 | 0.2 | < 0.3 | 0.13 | 0.14667 | 0.1 | 0.13 | 0.13 | 0.13
| Mean |  | 0.13167 |  |  | 0.12667 | 0.13033 | 0.11333 |  | 0.13833 | 0.16667 |  | 0.12667 | 0.17178 | 0.1 | 0.13333 | 0.13333 | 0.13667
| Median |  | 0.13 |  |  | 0.13 | 0.131 | 0.115 |  | 0.14 | 0.2 |  | 0.13 | 0.15737 | 0.1 | 0.13 | 0.13 | 0.135
| Std Dev. |  | 0.0040825 |  |  | 0.005164 | 0.012956 | 0.008165 |  | 0.0075277 | 0.05164 |  | 0.008165 | 0.037719 | 1.5202e-17 | 0.005164 | 0.005164 | 0.008165
| Rel.Std.Dev. |  | 3.101% |  |  | 4.077% | 9.941% | 7.204% |  | 5.442% | 30.98% |  | 6.446% | 21.96% | 1.52e-14% | 3.873% | 3.873% | 5.974%
| PDM3 |  | -0.3469% |  |  | -4.131% | -1.356% | -14.22% |  | 4.699% | 26.14% |  | -4.131% | 30.02% | -24.31% | 0.9145% | 0.9145% | 3.437%
| Z-Score (Absolute) |  | 0.17 | 3.37 | 492.25 | 0.67 | 0.3 | 2.02 | 11.8 | 0.51 | NA | 1.69 | 0.67 | 3.89 | NA | 0 | 0 | 0.34
Results from laboratories 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 262b (Certified Value 0.922 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.8747 | 1.2273 | 0.83188 | 0.88 | 1 | 1.21 | 0.97 | 0.96 | 0.75 | 0.85 | 0.69 | 1.064 | 1.25 | 0.81 | 0.73 | 0.77 | 0.923 |  |  | 1
| 1 | 2 | 0.8615 | 1.2924 | 0.85072 | 0.96 | 1.05 | 1.02 | 0.93 | 0.89 | 0.75 | 0.76 | 0.68 | 1.062 | 1.23 | 0.81 | 0.71 | 0.76 | 0.943 |  |  | 0
| 1 | 3 | 0.828 | 1.2009 | 0.83987 | 0.77 | 1.02 | 1.14 | 0.92 | 0.96 | 0.8 | 0.77 | 0.68 | 1.0691 | 1.23 | 0.81 | 0.74 | 0.76 | 0.882 |  |  | 16
| 1 | 4 | 0.8381 | 1.1663 | 0.84761 | 1.03 | 1.12 | 1.1 | 0.94 | 0.95 | 0.78 | 0.79 | 0.7 | 1.1123 | 1.24 | 0.79 | 0.76 | 0.76 | 0.908 |  |  | 0.92235
| 1 | 5 | 0.8633 | 1.2525 | 0.85224 | 0.89 | 1.04 | 1.22 | 0.98 | 0.95 | 0.77 | 0.84 | 0.68 | 1.1087 | 1.24 | 0.84 | 0.74 | 0.77 | 0.857 |  |  | 75
| 1 | 6 | 0.8476 | 1.1518 | 0.85332 | 0.8 | 1.07 | 1.14 | 0.95 | 0.95 | 0.72 | 0.77 | 0.7 | 1.1284 | 1.23 | 0.85 | 0.72 | 0.74 | 0.898
| Mean |  | 0.8522 | 1.2152 | 0.84594 | 0.88833 | 1.05 | 1.1383 | 0.94833 | 0.94333 | 0.76167 | 0.79667 | 0.68833 | 1.0907 | 1.2367 | 0.81833 | 0.73333 | 0.76 | 0.90183
| Median |  | 0.85455 | 1.2141 | 0.84916 | 0.885 | 1.045 | 1.14 | 0.945 | 0.95 | 0.76 | 0.78 | 0.685 | 1.0889 | 1.235 | 0.81 | 0.735 | 0.76 | 0.903
| Std Dev. |  | 0.017445 | 0.053134 | 0.0084181 | 0.09704 | 0.041952 | 0.073869 | 0.023166 | 0.026583 | 0.027869 | 0.038816 | 0.0098319 | 0.029031 | 0.008165 | 0.022286 | 0.017512 | 0.010954 | 0.030328
| Rel.Std.Dev. |  | 2.047% | 4.372% | 0.9951% | 10.92% | 3.995% | 6.489% | 2.443% | 2.818% | 3.659% | 4.872% | 1.428% | 2.662% | 0.6602% | 2.723% | 2.388% | 1.441% | 3.363%
| PDM3 |  | -7.605% | 31.75% | -8.284% | -3.688% | 13.84% | 23.42% | 2.817% | 2.275% | -17.42% | -13.63% | -25.37% | 18.26% | 34.08% | -11.28% | -20.49% | -17.6% | -2.224%
| Z-Score (Absolute) |  | 0.19 | 1.74 | 0.23 | 0 | 0.86 | 1.33 | 0.32 | 0.29 | 0.67 | 0.49 | 1.06 | 1.08 | 1.85 | 0.37 | 0.83 | 0.68 | 0.07
| Analytical results for As in OREAS 262b (Certified Value 13.9 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 15 | 14.423 | < 100 | 14 | 13 | 14.3 | 15.2 | 14.38 | 13.8 | 14 | 11.958 | 12.4 | 12.8442 | 14.1 | 12.7 | 13.8 | 13.4 | 12 |  | 1
| 1 | 2 | 14 | 14.6101 | < 100 | 17 | 13 | 12.6 | 15 | 14.105 | 14 | 12.7 | 11.652 | 12.6 | 13.0986 | 14.5 | 13.3 | 13.8 | 13 | 12 |  | 22
| 1 | 3 | 14 | 14.3573 | < 100 | 14 | 13 | 12.9 | 14.9 | 14.25 | 14.5 | 13.6 | 11.819 | 12.4 | 14.2552 | 14.2 | 12.8 | 14.4 | 13.9 | 12 |  | 16
| 1 | 4 | 14 | 14.121 | < 100 | 16 | 13 | 13.6 | 15.4 | 14.44 | 13.8 | 14.6 | 11.821 | 13.2 | 14.318 | 14.4 | 13 | 15 | 13.4 | 12 |  | 13.8769
| 1 | 5 | 14 | 15.1539 | < 100 | 17 | 13 | 13.9 | 15.4 | 14.12 | 13.8 | 13.6 | 12.017 | 12.9 | 13.5902 | 14.5 | 13.2 | 13.8 | 13.2 | 12 |  | 76
| 1 | 6 | 14 | 14.2645 | < 100 | 13 | 13 | 13.8 | 14.8 | 14.2467 | 13.6 | 14.1 | 12.848 | 13.4 | 13.7174 | 14 | 13.4 | 14.6 | 12.9 | 12
| Mean |  | 14.1667 | 14.4883 |  | 15.1667 | 13 | 13.5167 | 15.1167 | 14.2569 | 13.9167 | 13.7667 | 12.0192 | 12.8167 | 13.6372 | 14.2833 | 13.0667 | 14.2333 | 13.3 | 12
| Median |  | 14 | 14.3902 |  | 15 | 13 | 13.7 | 15.1 | 14.2483 | 13.8 | 13.8 | 11.8895 | 12.75 | 13.6538 | 14.3 | 13.1 | 14.1 | 13.3 | 12
| Std Dev. |  | 0.40825 | 0.36448 |  | 1.7224 | 0 | 0.64317 | 0.25626 | 0.13465 | 0.31252 | 0.64083 | 0.42541 | 0.42151 | 0.59549 | 0.2137 | 0.28048 | 0.51251 | 0.35777 | 0
| Rel.Std.Dev. |  | 2.882% | 2.516% |  | 11.36% | 0% | 4.758% | 1.695% | 0.9444% | 2.246% | 4.655% | 3.539% | 3.289% | 4.367% | 1.496% | 2.146% | 3.601% | 2.69% | 0%
| PDM3 |  | 2.088% | 4.406% |  | 9.295% | -6.319% | -2.596% | 8.934% | 2.739% | 0.2868% | -0.7942% | -13.39% | -7.64% | -1.727% | 2.929% | -5.839% | 2.569% | -4.157% | -13.53%
| Z-Score (Absolute) |  | NA | 0.67 | 42.57 | 1.47 | NA | 0.47 | 1.42 | 0.4 | 0 | 0.18 | 2.24 | 1.3 | 0.33 | 0.43 | 1 | 0.37 | 0.73 | NA
Results from laboratories 1.01, 1.07 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 262b (Certified Value < 10 ppm) | [20 empty] | Certified
B | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.06; 1.05: 1.09; 1.06: 1.12; 1.07: 1.13; 1.09: 1.15; 1.10: 1.16; 1.11: 1.19; 1.12: 1.20; 1.13: 1.21 |  | ppm
|  |  | 1.01: AR*OES; 1.03: AR*OES; 1.05: AR*OES/MS; 1.06: AR*MS; 1.07: AR*MS; 1.09: AR*MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS |  | 1
|  |  | 1.01: 25g; 1.03: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.09: 0.5g; 1.10: 01g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g |  | 1
| 1 | 1 | 1.01: < 10; 1.03: 1.7; 1.05: 1; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 1
| 1 | 2 | 1.01: < 10; 1.03: 0.3; 1.05: 1; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 0
| 1 | 3 | 1.01: < 10; 1.03: 2.2; 1.05: 2; 1.06: < 10; 1.07: 23; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 16
| 1 | 4 | 1.01: < 10; 1.03: 0.6; 1.05: 2; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | < 10
| 1 | 5 | 1.01: < 10; 1.03: 0.8; 1.05: 2; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 77
| 1 | 6 | 1.01: < 10; 1.03: 2.5; 1.05: 1; 1.06: < 10; 1.07: 64; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10
| Mean |  | 1.03: 1.35; 1.05: 1.5; 1.07: 43.5
| Median |  | 1.03: 1.25; 1.05: 1.5; 1.07: 43.5
| Std Dev. |  | 1.03: 0.9094; 1.05: 0.54772; 1.07: 28.9914
| Rel.Std.Dev. |  | 1.03: 67.36%; 1.05: 36.51%; 1.07: 66.65%
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.13
| Analytical results for Ba in OREAS 262b (Certified Value 75 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 70 | 87.7608 | 71.1157 | 71.646 | 90 | 74 | 80.3 | 76 | 74.3733 | 70 | 67 | 70 | 78.4687 | 70 | 70 | 70 | 76 |  |  | 1
| 1 | 2 | 67 | 91.8324 | 71.6811 | 75.66 | 93 | 75 | 74.5 | 76 | 71.425 | 70 | 66 | 70 | 77.6263 | 70 | 70 | 70 | 82 |  |  | 23
| 1 | 3 | 70 | 91.0548 | 71.6512 | 73.896 | 84 | 75 | 78.2 | 75 | 74.29 | 70 | 69 | 70 | 75.8911 | 70 | 70 | 70 | 79 |  |  | 16
| 1 | 4 | 67 | 85.968 | 71.3635 | 75.712 | 98 | 76 | 75.7 | 76 | 74.6033 | 70 | 72 | 70 | 84.6771 | 70 | 70 | 70 | 77 |  |  | 74.5221
| 1 | 5 | 70 | 89.5212 | 73.1978 | 73.775 | 88 | 74 | 81.2 | 77 | 71.64 | 70 | 69 | 70 | 83.4497 | 80 | 70 | 70 | 78 |  |  | 78
| 1 | 6 | 69 | 85.2444 | 71.6721 | 71.652 | 89 | 75 | 77.4 | 77 | 73.5867 | 70 | 75 | 70 | 82.9731 | 80 | 70 | 70 | 79
| Mean |  | 68.8333 | 88.5636 | 71.7802 | 73.7235 | 90.3333 | 74.8333 | 77.8833 | 76.1667 | 73.3197 | 70 | 69.6667 | 70 | 80.5144 | 73.3333 | 70 | 70 | 78.5
| Median |  | 69.5 | 88.641 | 71.6617 | 73.8355 | 89.5 | 75 | 77.8 | 76 | 73.9383 | 70 | 69 | 70 | 80.7209 | 70 | 70 | 70 | 78.5
| Std Dev. |  | 1.472 | 2.691 | 0.7298 | 1.808 | 4.761 | 0.75277 | 2.5841 | 0.75277 | 1.427 | 0 | 3.3267 | 0 | 3.6301 | 5.164 | 0 | 0 | 2.0736
| Rel.Std.Dev. |  | 2.138% | 3.038% | 1.017% | 2.452% | 5.27% | 1.006% | 3.318% | 0.9883% | 1.946% | 0% | 4.775% | 0% | 4.509% | 7.042% | 0% | 0% | 2.642%
| PDM3 |  | -7.634% | 18.84% | -3.679% | -1.072% | 21.22% | 0.4176% | 4.51% | 2.207% | -1.613% | -6.068% | -6.515% | -6.068% | 8.041% | -1.595% | -6.068% | -6.068% | 5.338%
| Z-Score (Absolute) |  | 1.34 | 2.62 | 0.75 | 0.36 | 2.98 | 0.13 | 0.48 | 0.13 | 0.44 | NA | 1.17 | NA | 1.01 | NA | NA | NA | 0.6
Results from laboratories 1.12, 1.15, 1.19, 1.20 and 1.21 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 262b (Certified Value 0.47 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 0.54 | < 0.5 | < 10 | 0.9 | < 0.5 | 0.5 | 0.46 | < 5 | 0.44 | 0.41 | 0.50053 | 0.6 | 0.45 | 0.43 | 0.44 | 0.59 |  |  |  | 1
| 1 | 2 | 0.49 | < 0.5 | < 10 | 0.9 | < 0.5 | 0.5 | 0.46 | < 5 | 0.44 | 0.41 | < 0.5 | < 0.5 | 0.45 | 0.42 | 0.43 | 0.56 |  |  |  | 24
| 1 | 3 | 0.51 | < 0.5 | < 10 | 0.9 | < 0.5 | 0.5 | 0.44 | < 5 | 0.46 | 0.41 | < 0.5 | < 0.5 | 0.45 | 0.43 | 0.43 | 0.57 |  |  |  | 16
| 1 | 4 | 0.49 | < 0.5 | < 10 | 0.9 | < 0.5 | 0.5 | 0.45 | < 5 | 0.45 | 0.42 | 0.50136 | 0.5 | 0.44 | 0.44 | 0.43 | 0.57 |  |  |  | 0.47075
| 1 | 5 | 0.5 | 0.50709 | < 10 | 0.8 | < 0.5 | 0.6 | 0.44 | < 5 | 0.44 | 0.41 | < 0.5 | 0.5 | 0.47 | 0.43 | 0.44 | 0.58 |  |  |  | 79
| 1 | 6 | 0.5 | < 0.5 | < 10 | 0.9 | < 0.5 | 0.4 | 0.47 | < 5 | 0.42 | 0.42 | 0.50044 | < 0.5 | 0.48 | 0.43 | 0.42 | 0.58
| Mean |  | 0.505 | 0.50709 |  | 0.88333 |  | 0.5 | 0.45333 |  | 0.44167 | 0.41333 | 0.50078 | 0.53333 | 0.45667 | 0.43 | 0.43167 | 0.575
| Median |  | 0.5 | 0.50709 |  | 0.9 |  | 0.5 | 0.455 |  | 0.44 | 0.41 | 0.50053 | 0.5 | 0.45 | 0.43 | 0.43 | 0.575
| Std Dev. |  | 0.018708 |  |  | 0.040825 |  | 0.063246 | 0.012111 |  | 0.013292 | 0.005164 | 0.00050775 | 0.057735 | 0.015055 | 0.0063246 | 0.0075277 | 0.010488
| Rel.Std.Dev. |  | 3.705% |  |  | 4.622% |  | 12.65% | 2.671% |  | 3.009% | 1.249% | 0.1014% | 10.83% | 3.297% | 1.471% | 1.744% | 1.824%
| PDM3 |  | 7.275% | 7.718% |  | 87.64% |  | 6.213% | -3.7% |  | -6.179% | -12.2% | 6.378% | 13.29% | -2.992% | -8.657% | -8.303% | 22.14%
| Z-Score (Absolute) |  | 0.67 | 1.58 | 48.53 | NA | 2.04 | NA | 0.12 | 21.92 | 0 | 0.3 | 0.71 | NA | 0.16 | 0.12 | 0.11 | 1.42
Results from laboratories 1.06, 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 262b (Certified Value 0.69 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.73 | 0.65122 | < 10 | < 0.3 | 0.75 | 0.7 | 0.67 | < 2 | 0.7 | 0.7 | 0.667 | 0.6 | 0.74419 | 0.78 | 0.69 | 0.71 | 0.63 | 0.86 |  | 1
| 1 | 2 | 0.69 | 0.66921 | < 10 | < 0.3 | 0.71 | 0.65 | 0.65 | < 2 | 0.7 | 0.7 | 0.626 | 0.63 | 0.67132 | 0.77 | 0.76 | 0.71 | 0.66 | 0.86 |  | 25
| 1 | 3 | 0.71 | 0.65759 | < 10 | < 0.3 | 0.72 | 0.67 | 0.63 | < 2 | 0.68 | 0.7 | 0.652 | 0.62 | 0.64758 | 0.77 | 0.7 | 0.7 | 0.67 | 0.89 |  | 16
| 1 | 4 | 0.72 | 0.6653 | < 10 | < 0.3 | 0.72 | 0.64 | 0.64 | < 2 | 0.67 | 0.6 | 0.64 | 0.62 | 0.74163 | 0.81 | 0.7 | 0.69 | 0.67 | 0.85 |  | 0.68632
| 1 | 5 | 0.71 | 0.67491 | < 10 | < 0.3 | 0.72 | 0.71 | 0.96 | < 2 | 0.68 | 0.7 | 0.673 | 0.64 | 0.7099 | 0.77 | 0.74 | 0.74 | 0.63 | 0.81 |  | 80
| 1 | 6 | 0.71 | 0.7095 | < 10 | < 0.3 | 0.71 | 0.69 | 0.64 | < 2 | 0.64 | 0.7 | 0.67 | 0.62 | 0.68431 | 0.75 | 0.73 | 0.72 | 0.63 | 0.93
| Mean |  | 0.71167 | 0.67129 |  |  | 0.72167 | 0.67667 | 0.69833 |  | 0.67833 | 0.68333 | 0.65467 | 0.62167 | 0.69982 | 0.775 | 0.72 | 0.71167 | 0.64833 | 0.86667
| Median |  | 0.71 | 0.66726 |  |  | 0.72 | 0.68 | 0.645 |  | 0.68 | 0.7 | 0.6595 | 0.62 | 0.69711 | 0.77 | 0.715 | 0.71 | 0.645 | 0.86
| Std Dev. |  | 0.013292 | 0.020514 |  |  | 0.01472 | 0.028048 | 0.12891 |  | 0.022286 | 0.040825 | 0.018801 | 0.013292 | 0.038988 | 0.019748 | 0.027568 | 0.017224 | 0.020412 | 0.040332
| Rel.Std.Dev. |  | 1.868% | 3.056% |  |  | 2.04% | 4.145% | 18.46% |  | 3.285% | 5.974% | 2.872% | 2.138% | 5.571% | 2.548% | 3.829% | 2.42% | 3.148% | 4.654%
| PDM3 |  | 3.693% | -2.19% |  |  | 5.15% | -1.406% | 1.751% |  | -1.163% | -0.4349% | -4.612% | -9.42% | 1.968% | 12.92% | 4.908% | 3.693% | -5.535% | 26.28%
| Z-Score (Absolute) |  | 0.28 | 0.67 | 101.62 | 12.99 | 0.52 | 0.55 | 0.04 | 7.09 | 0.51 | NA | 1.07 | 1.85 | 0 | 1.78 | 0.48 | 0.28 | 1.22 | 3.94
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 262b (Certified Value 0.266 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 0.28 | 0.29563 | 0.26236 | 0.27702 | 0.94 | 0.27 | 0.28 | 0.25 | 0.27 | 0.27 | 0.29 | 0.25276 | 0.26 | 0.25431 | 0.27 | 0.28 | 0.27 | 0.27 | 0.257 | 1
| 1 | 2 | 0.28 | 0.29323 | 0.26616 | 0.27651 | 0.98 | 0.28 | 0.25 | 0.25 | 0.27 | 0.27 | 0.27 | 0.25146 | 0.25 | 0.26498 | 0.26 | 0.28 | 0.26 | 0.27 | 0.263 | 0
| 1 | 3 | 0.26 | 0.30836 | 0.26203 | 0.2753 | 0.94 | 0.28 | 0.26 | 0.25 | 0.27 | 0.28 | 0.24 | 0.25346 | 0.26 | 0.25991 | 0.26 | 0.28 | 0.26 | 0.27 | 0.262 | 16
| 1 | 4 | 0.28 | 0.30375 | 0.2629 | 0.27827 | 1.02 | 0.28 | 0.25 | 0.25 | 0.27 | 0.28 | 0.25 | 0.25586 | 0.26 | 0.26527 | 0.26 | 0.28 | 0.27 | 0.27 | 0.25 | 0.26614
| 1 | 5 | 0.27 | 0.30885 | 0.26771 | 0.27988 | 0.93 | 0.28 | 0.27 | 0.25 | 0.27 | 0.27 | 0.28 | 0.25996 | 0.26 | 0.25521 | 0.27 | 0.29 | 0.26 | 0.27 | 0.259 | 81
| 1 | 6 | 0.27 | 0.29929 | 0.26631 | 0.27803 | 1.03 | 0.28 | 0.26 | 0.25 | 0.28 | 0.27 | 0.25 | 0.25816 | 0.26 | 0.27045 | 0.26 | 0.3 | 0.26 | 0.27 | 0.249
| Mean |  | 0.27333 | 0.30152 | 0.26458 | 0.2775 | 0.97333 | 0.27833 | 0.26167 | 0.25 | 0.27167 | 0.27333 | 0.26333 | 0.25528 | 0.25833 | 0.26169 | 0.26333 | 0.285 | 0.26333 | 0.27 | 0.25667
| Median |  | 0.275 | 0.30152 | 0.26453 | 0.27753 | 0.96 | 0.28 | 0.26 | 0.25 | 0.27 | 0.27 | 0.26 | 0.25466 | 0.26 | 0.26244 | 0.26 | 0.28 | 0.26 | 0.27 | 0.258
| Std Dev. |  | 0.008165 | 0.0065398 | 0.0024312 | 0.0015898 | 0.043665 | 0.0040825 | 0.01169 | 0 | 0.0040825 | 0.005164 | 0.019664 | 0.0033088 | 0.0040825 | 0.0063262 | 0.005164 | 0.0083666 | 0.005164 | 0 | 0.0059554
| Rel.Std.Dev. |  | 2.987% | 2.169% | 0.9189% | 0.5729% | 4.486% | 1.467% | 4.468% | 0% | 1.503% | 1.889% | 7.467% | 1.296% | 1.58% | 2.417% | 1.961% | 2.936% | 1.961% | 0% | 2.32%
| PDM3 |  | 2.703% | 13.29% | -0.5863% | 4.27% | 265.7% | 4.582% | -1.681% | -6.064% | 2.077% | 2.703% | -1.055% | -4.082% | -2.933% | -1.673% | -1.055% | 7.086% | -1.055% | 1.45% | -3.56%
| Z-Score (Absolute) |  | 0.75 | 3.15 | 0 | 1.1 | 60.4 | 1.17 | 0.25 | 1.24 | 0.6 | 0.75 | 0.11 | 0.79 | 0.53 | 0.25 | 0.11 | 1.74 | 0.11 | 0.46 | 0.67
| Analytical results for Cd in OREAS 262b (Certified Value 0.063 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.09 | 0.14 | < 0.5 | 0.06 | 0.1 | 0.06 | 0.082633 | 0.06 | 0.05 | 0.07 | 0.06 |  |  |  | 1
| 1 | 2 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.09 | 0.14 | < 0.5 | 0.07 | 0.1 | 0.05 | 0.075832 | 0.06 | 0.06 | 0.05 | 0.06 |  |  |  | 26
| 1 | 3 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.07 | 0.14 | < 0.5 | 0.06 | 0.1 | 0.06 | 0.07797 | 0.06 | 0.05 | 0.06 | 0.05 |  |  |  | 16
| 1 | 4 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.03 | 0.14 | < 0.5 | 0.06 | 0.1 | 0.07 | 0.10837 | 0.06 | 0.06 | 0.06 | 0.05 |  |  |  | 0.062926
| 1 | 5 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.06 | 0.15 | < 0.5 | 0.06 | 0.1 | 0.06 | 0.087478 | 0.07 | 0.06 | 0.07 | 0.06 |  |  |  | 82
| 1 | 6 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.07 | 0.12 | 0.15 | < 0.5 | 0.05 | 0.1 | 0.07 | 0.1059 | 0.06 | 0.06 | 0.08 | 0.06
| Mean |  | 0.065 |  |  |  | 0.071667 | 0.076667 | 0.14333 |  | 0.06 | 0.1 | 0.061667 | 0.089698 | 0.061667 | 0.056667 | 0.065 | 0.056667
| Median |  | 0.065 |  |  |  | 0.07 | 0.08 | 0.14 |  | 0.06 | 0.1 | 0.06 | 0.085056 | 0.06 | 0.06 | 0.065 | 0.06
| Std Dev. |  | 0.0054772 |  |  |  | 0.0040825 | 0.030768 | 0.005164 |  | 0.0063246 | 1.5202e-17 | 0.0075277 | 0.014113 | 0.0040825 | 0.005164 | 0.010488 | 0.005164
| Rel.Std.Dev. |  | 8.427% |  |  |  | 5.696% | 40.13% | 3.603% |  | 10.54% | 1.52e-14% | 12.21% | 15.73% | 6.62% | 9.113% | 16.14% | 9.113%
| PDM3 |  | 3.296% |  |  |  | 13.89% | 21.84% | 127.8% |  | -4.65% | 58.92% | -2.001% | 42.55% | -2.001% | -9.947% | 3.296% | -9.947%
| Z-Score (Absolute) |  | 0.3 | 1.27 | 221.55 | 3.52 | 0 | 0.22 | 3.22 | 8.02 | 0.52 | NA | 0.45 | 0.81 | 0.45 | 0.67 | 0.3 | 0.67
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 262b (Certified Value 31.3 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 33.391 | 30.2068 | 33.4 | 34.9 | 33.32 | 29.8 | 27.6 | 32.5772 | 29.3 | 30.5 | 35.5 | 28.6 | 31.99 | [6 empty] | 1
| 1 | 2 | 32.317 | 30.5087 | 33.9 | 31.4 | 32.65 | 30 | 27.9 | 33.322 | 28.3 | 30.9 | 31.5 | 29 | 31.53 | [6 empty] | 27
| 1 | 3 | 33.402 | 30.0036 | 33.8 | 32.6 | 32.12 | 31.1 | 27.4 | 32.7558 | 28.9 | 30.1 | 30.4 | 29.1 | 31.81 | [6 empty] | 16
| 1 | 4 | 31.633 | 29.5807 | 33.2 | 31.6 | 32.6 | 30.3 | 29 | 34.3121 | 28.6 | 29.7 | 33.5 | 29.6 | 31.55 | [6 empty] | 31.2757
| 1 | 5 | 32.514 | 30.7787 | 33.2 | 33.2 | 33.37 | 30.2 | 28.1 | 33.2203 | 29.7 | 31.3 | 32 | 27.9 | 30.06 | [6 empty] | 83
| 1 | 6 | 31.773 | 30.077 | 33.5 | 33.5 | 32.93 | 29.7 | 29.7 | 34.2624 | 28.1 | 32.3 | 34.7 | 26.9 | 31.85
| Mean |  | 32.505 | 30.1926 | 33.5 | 32.8667 | 32.8317 | 30.1833 | 28.2833 | 33.4083 | 28.8167 | 30.8 | 32.9333 | 28.5167 | 31.465
| Median |  | 32.4155 | 30.1419 | 33.45 | 32.9 | 32.79 | 30.1 | 28 | 33.2711 | 28.75 | 30.7 | 32.75 | 28.8 | 31.68
| Std Dev. |  | 0.76436 | 0.41647 | 0.29665 | 1.3018 | 0.47562 | 0.50365 | 0.88863 | 0.73559 | 0.608 | 0.92736 | 1.9684 | 0.97451 | 0.7111
| Rel.Std.Dev. |  | 2.352% | 1.379% | 0.8855% | 3.961% | 1.449% | 1.669% | 3.142% | 2.202% | 2.11% | 3.011% | 5.977% | 3.417% | 2.26%
| PDM3 |  | 3.931% | -3.463% | 7.112% | 5.087% | 4.975% | -3.493% | -9.568% | 6.819% | -7.862% | -1.521% | 5.3% | -8.822% | 0.6054%
| Z-Score (Absolute) |  | 0.5 | 0.61 | 0.98 | 0.67 | 0.66 | 0.62 | 1.53 | 0.93 | 1.27 | 0.32 | 0.71 | 1.42 | 0
| Analytical results for Co in OREAS 262b (Certified Value 4.02 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 3.9 | < 5 | 3.8483 | < 10 | < 1 | 4.3 | 4.9 | 4.4 | 4.1533 | 3.4 | 3.8 | 3.3 | 4.2863 | 4.51 | 3.8 | 4 | 3.7 | 4.1 |  | 1
| 1 | 2 | 3.8 | < 5 | 3.8859 | < 10 | < 1 | 4.4 | 4.4 | 4.2 | 3.96 | 3.4 | 3.7 | 3.3 | 4.2316 | 4.55 | 4 | 3.7 | 3.7 | 4.2 |  | 28
| 1 | 3 | 3.9 | < 5 | 3.8769 | 10.866 | < 1 | 4.3 | 4.7 | 4.1 | 4.14 | 3.5 | 3.6 | 3.1 | 4.2639 | 4.55 | 3.9 | 3.6 | 3.7 | 4.2 |  | 16
| 1 | 4 | 3.8 | < 5 | 3.8051 | < 10 | 58 | 4.4 | 4.8 | 4.2 | 4.2 | 3.4 | 3.6 | 3.4 | 4.9458 | 4.61 | 3.9 | 4 | 3.7 | 4.2 |  | 4.0244
| 1 | 5 | 3.8 | < 5 | 4.2371 | 11.606 | < 1 | 4.2 | 5 | 4.4 | 4.16 | 3.4 | 3.9 | 3.2 | 4.4139 | 4.6 | 4.1 | 3.7 | 3.6 | 4.2 |  | 84
| 1 | 6 | 3.9 | < 5 | 3.826 | 10.793 | < 1 | 4.3 | 4.9 | 4.2 | 4.12 | 3.6 | 3.8 | 3.3 | 4.9516 | 4.46 | 4.1 | 4 | 3.6 | 4.2
| Mean |  | 3.85 |  | 3.9132 | 11.0883 | 58 | 4.3167 | 4.7833 | 4.25 | 4.1222 | 3.45 | 3.7333 | 3.2667 | 4.5155 | 4.5467 | 3.9667 | 3.8333 | 3.6667 | 4.1833
| Median |  | 3.85 |  | 3.8626 | 10.866 | 58 | 4.3 | 4.85 | 4.2 | 4.1467 | 3.4 | 3.75 | 3.3 | 4.3501 | 4.55 | 3.95 | 3.85 | 3.7 | 4.2
| Std Dev. |  | 0.054772 |  | 0.16153 | 0.4498 |  | 0.075277 | 0.2137 | 0.12247 | 0.083763 | 0.083666 | 0.12111 | 0.10328 | 0.34122 | 0.056095 | 0.12111 | 0.18619 | 0.05164 | 0.040825
| Rel.Std.Dev. |  | 1.423% |  | 4.128% | 4.056% |  | 1.744% | 4.468% | 2.882% | 2.032% | 2.425% | 3.244% | 3.162% | 7.557% | 1.234% | 3.053% | 4.857% | 1.408% | 0.9759%
| PDM3 |  | -4.332% |  | -2.761% | 175.5% | 1341% | 7.264% | 18.86% | 5.607% | 2.432% | -14.27% | -7.231% | -18.83% | 12.2% | 12.98% | -1.433% | -4.747% | -8.888% | 3.95%
| Z-Score (Absolute) |  | 0.38 | 3.02 | 0.26 | 7.83 | 11.82 | 0.53 | 1.45 | 0.4 | 0.15 | 1.16 | 0.61 | 1.52 | 0.92 | 0.98 | 0.15 | 0.41 | 0.74 | 0.27
| Analytical results for Cr in OREAS 262b (Certified Value 31.8 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 34 | 33.1668 | 31.4782 | 29.534 | 32 | 35 | 32 | 32 | 34.2295 | 32 | 34 | 33 | 30.267 | 32 | 32 | 32 | 31 | 31 |  | 1
| 1 | 2 | 33 | 33.9984 | 31.54 | 29.672 | 33 | 35 | 29 | 32 | 33.7121 | 32 | 31 | 30 | 30.2213 | 32 | 32 | 31 | 31 | 32 |  | 29
| 1 | 3 | 31 | 34.4628 | 31.3655 | 29.969 | 33 | 36 | 31 | 32 | 33.8615 | 32 | 29 | 30 | 30.8471 | 32 | 33 | 32 | 31 | 31 |  | 16
| 1 | 4 | 33 | 32.616 | 31.1119 | 30.258 | 34 | 36 | 30 | 32 | 33.835 | 32 | 29 | 30 | 32.0714 | 34 | 32 | 32 | 31 | 31 |  | 31.8
| 1 | 5 | 33 | 34.1064 | 32.2686 | 30.125 | 35 | 36 | 32 | 33 | 33.5953 | 31 | 32 | 30 | 32.2833 | 33 | 33 | 32 | 32 | 29 |  | 85
| 1 | 6 | 32 | 34.454 | 31.7927 | 29.236 | 31 | 36 | 31 | 32 | 33.8932 | 32 | 30 | 30 | 31.6302 | 32 | 35 | 32 | 31 | 30
| Mean |  | 32.6667 | 33.8007 | 31.5928 | 29.799 | 33 | 35.6667 | 30.8333 | 32.1667 | 33.8544 | 31.8333 | 30.8333 | 30.5 | 31.2201 | 32.5 | 32.8333 | 31.8333 | 31.1667 | 30.6667
| Median |  | 33 | 34.0524 | 31.5091 | 29.8205 | 33 | 36 | 31 | 32 | 33.8482 | 32 | 30.5 | 30 | 31.2386 | 32 | 32.5 | 32 | 31 | 31
| Std Dev. |  | 1.0328 | 0.74874 | 0.39889 | 0.38709 | 1.4142 | 0.5164 | 1.169 | 0.40825 | 0.21446 | 0.40825 | 1.9408 | 1.2247 | 0.90204 | 0.83666 | 1.169 | 0.40825 | 0.40825 | 1.0328
| Rel.Std.Dev. |  | 3.162% | 2.215% | 1.263% | 1.299% | 4.285% | 1.448% | 3.791% | 1.269% | 0.6335% | 1.282% | 6.294% | 4.016% | 2.889% | 2.574% | 3.561% | 1.282% | 1.31% | 3.368%
| PDM3 |  | 2.725% | 6.292% | -0.6516% | -6.293% | 3.774% | 12.16% | -3.04% | 1.153% | 6.46% | 0.1048% | -3.04% | -4.088% | -1.824% | 2.201% | 3.249% | 0.1048% | -1.992% | -3.564%
| Z-Score (Absolute) |  | 0.56 | 1.33 | 0.16 | 1.37 | 0.79 | 2.58 | 0.67 | 0.22 | 1.36 | 0 | 0.67 | 0.9 | 0.41 | 0.45 | 0.67 | 0 | 0.45 | 0.79
| Analytical results for Cs in OREAS 262b (Certified Value 1.65 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 1.81 | 1.6461 | 2.08 | 1.63 | 1.82 | 1.38 | 1.33 | 1.6495 | 2.26 | 1.39 | 1.57 | 1.39 | 1.67 | [6 empty] | 1
| 1 | 2 | 1.75 | 1.6253 | 2.02 | 1.41 | 1.84 | 1.39 | 1.27 | 1.6957 | 2.27 | 1.47 | 1.45 | 1.38 | 1.67 | [6 empty] | 30
| 1 | 3 | 1.79 | 1.6403 | 2.09 | 1.59 | 1.78 | 1.49 | 1.25 | 1.6883 | 2.26 | 1.4 | 1.45 | 1.38 | 1.63 | [6 empty] | 16
| 1 | 4 | 1.71 | 1.6229 | 2.12 | 1.56 | 1.88 | 1.41 | 1.33 | 1.6789 | 2.3 | 1.41 | 1.5 | 1.4 | 1.64 | [6 empty] | 1.6534
| 1 | 5 | 1.8 | 1.6277 | 2.03 | 1.59 | 1.87 | 1.42 | 1.28 | 1.6501 | 2.29 | 1.48 | 1.48 | 1.36 | 1.57 | [6 empty] | 86
| 1 | 6 | 1.77 | 1.6608 | 2.07 | 1.61 | 1.85 | 1.43 | 1.34 | 1.7013 | 2.19 | 1.47 | 1.47 | 1.3 | 1.62
| Mean |  | 1.7717 | 1.6372 | 2.0683 | 1.565 | 1.84 | 1.42 | 1.3 | 1.6773 | 2.2617 | 1.4367 | 1.4867 | 1.3683 | 1.6333
| Median |  | 1.78 | 1.634 | 2.075 | 1.59 | 1.845 | 1.415 | 1.305 | 1.6836 | 2.265 | 1.44 | 1.475 | 1.38 | 1.635
| Std Dev. |  | 0.037103 | 0.0147 | 0.037639 | 0.079436 | 0.036332 | 0.038987 | 0.037947 | 0.022583 | 0.038687 | 0.040825 | 0.045019 | 0.036009 | 0.037238
| Rel.Std.Dev. |  | 2.094% | 0.8979% | 1.82% | 5.076% | 1.975% | 2.746% | 2.919% | 1.346% | 1.711% | 2.842% | 3.028% | 2.632% | 2.28%
| PDM3 |  | 7.153% | -0.9805% | 25.1% | -5.346% | 11.29% | -14.12% | -21.37% | 1.446% | 36.79% | -13.11% | -10.08% | -17.24% | -1.213%
| Z-Score (Absolute) |  | 0.47 | 0.01 | 1.49 | 0.23 | 0.71 | 0.73 | 1.14 | 0.15 | 2.15 | 0.67 | 0.5 | 0.91 | 0
| Analytical results for Cu in OREAS 262b (Certified Value 66 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 64 | 65.9576 | 67.7399 | 66.625 | 71 | 69 | 71.3 | 65.9 | 69.9233 | 65 | 70.2 | 63.7845 | 67.4 | 64.1798 | 64.3 | 64.7 | 68.2 | 66 | 60.2 | 1
| 1 | 2 | 62.4 | 67.2408 | 63.5471 | 67.105 | 67 | 68 | 68.7 | 63.5 | 66.585 | 64.8 | 65.3 | 63.0673 | 66.4 | 67.1864 | 64.2 | 66.9 | 64.4 | 63 | 60.3 | 31
| 1 | 3 | 62.2 | 66.8844 | 65.9647 | 67.122 | 67 | 69 | 68.4 | 64.6 | 69.5467 | 66.8 | 72.3 | 63.5402 | 67.6 | 63.5999 | 65.9 | 64.1 | 66 | 65.2 | 59.6 | 16
| 1 | 4 | 62.2 | 65.1888 | 66.8419 | 66.19 | 78 | 68 | 70.9 | 63.7 | 69.57 | 63.5 | 69.6 | 63.3698 | 68.9 | 66.6765 | 65.2 | 65.2 | 65.8 | 62.8 | 57 | 66.1824
| 1 | 5 | 63 | 65.666 | 67.2353 | 67.062 | 68 | 68 | 70.1 | 65.4 | 67.2 | 65.2 | 69.3 | 63.6422 | 67.2 | 64.6325 | 64.7 | 66.5 | 64 | 63 | 55.5 | 87
| 1 | 6 | 63.1 | 65.9232 | 65.0335 | 63.771 | 68 | 68 | 70.7 | 63.5 | 68.6267 | 63.2 | 72.2 | 63.0361 | 66.9 | 65.7628 | 62.6 | 67.2 | 65.2 | 62.6 | 57.8
| Mean |  | 62.8167 | 66.1435 | 66.0604 | 66.3125 | 69.8333 | 68.3333 | 70.0167 | 64.4333 | 68.5753 | 64.75 | 69.8167 | 63.4067 | 67.4 | 65.3397 | 64.4833 | 65.7667 | 65.6 | 63.7667 | 58.4
| Median |  | 62.7 | 65.9404 | 66.4033 | 66.8435 | 68 | 68 | 70.4 | 64.15 | 69.0867 | 64.9 | 69.9 | 63.455 | 67.3 | 65.1977 | 64.5 | 65.85 | 65.5 | 63 | 58.7
| Std Dev. |  | 0.69976 | 0.77143 | 1.5616 | 1.2975 | 4.2622 | 0.5164 | 1.204 | 1.0386 | 1.386 | 1.2988 | 2.5577 | 0.30656 | 0.84617 | 1.4317 | 1.1161 | 1.2738 | 1.4913 | 1.4501 | 1.9504
| Rel.Std.Dev. |  | 1.114% | 1.166% | 2.364% | 1.957% | 6.103% | 0.7557% | 1.72% | 1.612% | 2.021% | 2.006% | 3.663% | 0.4835% | 1.255% | 2.191% | 1.731% | 1.937% | 2.273% | 2.274% | 3.34%
| PDM3 |  | -5.086% | -0.05881% | -0.1843% | 0.1966% | 5.516% | 3.25% | 5.794% | -2.643% | 3.616% | -2.164% | 5.491% | -4.194% | 1.84% | -1.273% | -2.567% | -0.6281% | -0.88% | -3.65% | -11.76%
| Z-Score (Absolute) |  | 1.22 | 0.16 | 0.12 | 0.23 | 1.68 | 1.06 | 1.75 | 0.55 | 1.16 | 0.42 | 1.67 | 0.97 | 0.67 | 0.18 | 0.53 | 0 | 0.07 | 0.83 | 3.04
| Analytical results for Dy in OREAS 262b (Indicative Value 1.18 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 1.178; 1.02: 1.1715; 1.03: 1.3 | 1
| 1 | 2 | 1.01: 1.115; 1.02: 1.1399; 1.03: 1.2 | 5
| 1 | 3 | 1.01: 1.141; 1.02: 1.1685; 1.03: 1.3 | 16
| 1 | 4 | 1.01: 1.137; 1.02: 1.1069; 1.03: 1.1 | 1.1762
| 1 | 5 | 1.01: 1.128; 1.02: 1.2164; 1.03: 1.2 | 11
| 1 | 6 | 1.01: 1.16; 1.02: 1.2092; 1.03: 1.2
| Mean |  | 1.01: 1.1432; 1.02: 1.1687; 1.03: 1.2167
| Median |  | 1.01: 1.139; 1.02: 1.17; 1.03: 1.2
| Std Dev. |  | 1.01: 0.02264; 1.02: 0.041403; 1.03: 0.075277
| Rel.Std.Dev. |  | 1.01: 1.98%; 1.02: 3.543%; 1.03: 6.187%
| PDM3 |  | 1.01: -2.808%; 1.02: -0.6339%; 1.03: 3.441%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0; 1.03: 1.26
| Analytical results for Er in OREAS 262b (Indicative Value 0.5 ppm) | [20 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 0.516; 1.02: 0.49907; 1.03: 0.5 | 1
| 1 | 2 | 1.01: 0.497; 1.02: 0.46915; 1.03: 0.5 | 6
| 1 | 3 | 1.01: 0.51; 1.02: 0.49328; 1.03: 0.5 | 16
| 1 | 4 | 1.01: 0.519; 1.02: 0.48833; 1.03: 0.5 | 0.49935
| 1 | 5 | 1.01: 0.489; 1.02: 0.50969; 1.03: 0.5 | 12
| 1 | 6 | 1.01: 0.502; 1.02: 0.49573; 1.03: 0.5
| Mean |  | 1.01: 0.5055; 1.02: 0.49254; 1.03: 0.5
| Median |  | 1.01: 0.506; 1.02: 0.4945; 1.03: 0.5
| Std Dev. |  | 1.01: 0.011572; 1.02: 0.01351; 1.03: 0
| Rel.Std.Dev. |  | 1.01: 2.289%; 1.02: 2.743%; 1.03: 0%
| PDM3 |  | 1.01: 1.232%; 1.02: -1.363%; 1.03: 0.1306%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.91; 1.03: 0
| Analytical results for Eu in OREAS 262b (Indicative Value 0.43 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 0.429; 1.02: 0.448; 1.03: 0.5 | 1
| 1 | 2 | 1.01: 0.43; 1.02: 0.42294; 1.03: 0.4 | 7
| 1 | 3 | 1.01: 0.432; 1.02: 0.40916; 1.03: 0.4 | 16
| 1 | 4 | 1.01: 0.437; 1.02: 0.4345; 1.03: 0.4 | 0.43138
| 1 | 5 | 1.01: 0.433; 1.02: 0.43526; 1.03: 0.4 | 13
| 1 | 6 | 1.01: 0.427; 1.02: 0.42705; 1.03: 0.5
| Mean |  | 1.01: 0.43133; 1.02: 0.42948; 1.03: 0.43333
| Median |  | 1.01: 0.431; 1.02: 0.43078; 1.03: 0.4
| Std Dev. |  | 1.01: 0.0035024; 1.02: 0.013137; 1.03: 0.05164
| Rel.Std.Dev. |  | 1.01: 0.812%; 1.02: 3.059%; 1.03: 11.92%
| PDM3 |  | 1.01: -0.01166%; 1.02: -0.4403%; 1.03: 0.452%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 0.73
| Analytical results for Fe in OREAS 262b (Certified Value 2.07 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 2.16 | 2.2092 | 2.0348 | 2.0404 | 2.3 | 2.18 | 2.24 | 1.89 | 2.07 | 2.05 | 2.22 | 1.8696 | 1.92 | 2.0701 | 2.18 | 2.04 | 1.92 | 2 |  | 1
| 1 | 2 | 2.14 | 2.2662 | 2.0543 | 2.0415 | 2.27 | 2.14 | 2.24 | 1.86 | 2.06 | 2.07 | 2.08 | 1.898 | 1.91 | 2.0983 | 2.16 | 2.02 | 1.91 | 1.98 |  | 0
| 1 | 3 | 2 | 2.2649 | 2.029 | 2.0281 | 2.24 | 2.15 | 2.19 | 1.86 | 2.08 | 2.12 | 1.95 | 1.8794 | 1.93 | 2.0877 | 2.18 | 2.03 | 1.93 | 1.98 |  | 16
| 1 | 4 | 2.14 | 2.2115 | 2.0149 | 2.0522 | 2.34 | 2.16 | 2.25 | 1.88 | 2.07 | 2.09 | 1.98 | 1.8616 | 1.95 | 2.0731 | 2.18 | 2.03 | 1.96 | 1.98 |  | 2.0674
| 1 | 5 | 2.11 | 2.2518 | 2.092 | 2.0699 | 2.21 | 2.13 | 2.23 | 1.9 | 2.07 | 2.08 | 2.18 | 1.8693 | 1.9 | 2.1414 | 2.19 | 2.09 | 1.92 | 1.98 |  | 88
| 1 | 6 | 2.06 | 2.2057 | 2.0507 | 2.0544 | 2.28 | 2.19 | 2.27 | 1.86 | 2.09 | 2.03 | 2 | 1.8879 | 1.94 | 2.1362 | 2.13 | 2.13 | 1.93 | 1.96
| Mean |  | 2.1017 | 2.2349 | 2.046 | 2.0478 | 2.2733 | 2.1583 | 2.2367 | 1.875 | 2.0733 | 2.0733 | 2.0683 | 1.8777 | 1.925 | 2.1011 | 2.17 | 2.0567 | 1.9283 | 1.98
| Median |  | 2.125 | 2.2317 | 2.0427 | 2.0469 | 2.275 | 2.155 | 2.24 | 1.87 | 2.07 | 2.075 | 2.04 | 1.8745 | 1.925 | 2.093 | 2.18 | 2.035 | 1.925 | 1.98
| Std Dev. |  | 0.0608 | 0.02905 | 0.026783 | 0.01438 | 0.045461 | 0.023166 | 0.026583 | 0.017607 | 0.010328 | 0.031411 | 0.11143 | 0.013519 | 0.018708 | 0.030972 | 0.021909 | 0.043665 | 0.017224 | 0.012649
| Rel.Std.Dev. |  | 2.893% | 1.3% | 1.309% | 0.7022% | 2% | 1.073% | 1.189% | 0.939% | 0.4981% | 1.515% | 5.387% | 0.72% | 0.9719% | 1.474% | 1.01% | 2.123% | 0.8932% | 0.6388%
| PDM3 |  | 1.659% | 8.102% | -1.036% | -0.9482% | 9.962% | 4.4% | 8.189% | -9.305% | 0.2882% | 0.2882% | 0.04639% | -9.176% | -6.887% | 1.633% | 4.964% | -0.5179% | -6.725% | -4.226%
| Z-Score (Absolute) |  | 0.23 | 1.24 | 0.19 | 0.17 | 1.53 | 0.66 | 1.25 | 1.48 | 0.02 | 0.02 | 0.02 | 1.46 | 1.1 | 0.23 | 0.75 | 0.11 | 1.08 | 0.69
| Analytical results for Ga in OREAS 262b (Certified Value 2.93 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 2.58 | 2.8537 | 3.7 | 3.52 | 3.75 | 3.1693 | 2.29 | 2 | 2.21 | 2.8852 | 3.74 | 2.46 | 3 | 2.34 | [5 empty] | 1
| 1 | 2 | 2.51 | 2.5582 | 3.3 | 3.25 | 3.6 | 3.0867 | 2.32 | 2 | 2.19 | 2.9654 | 3.57 | 2.56 | 2.64 | 2.29 | [5 empty] | 35
| 1 | 3 | 2.48 | 2.8319 | 3.6 | 3.42 | 3.56 | 3.172 | 2.46 | 2 | 2.08 | 2.9416 | 3.59 | 2.48 | 2.64 | 2.34 | [5 empty] | 16
| 1 | 4 | 2.47 | 2.7522 | 3.7 | 3.2 | 3.67 | 3.2072 | 2.26 | 2 | 2.28 | 3.2736 | 3.65 | 2.45 | 2.84 | 2.3 | [5 empty] | 2.9263
| 1 | 5 | 2.54 | 3.0668 | 3.5 | 3.52 | 3.76 | 3.1437 | 2.27 | 2 | 2.16 | 3.2514 | 3.7 | 2.61 | 2.66 | 2.24 | [5 empty] | 89
| 1 | 6 | 2.57 | 2.9312 | 3.8 | 3.49 | 3.64 | 3.1757 | 2.37 | 2 | 2.2 | 3.2279 | 3.63 | 2.6 | 2.98 | 2.23
| Mean |  | 2.525 | 2.8323 | 3.6 | 3.4 | 3.6633 | 3.1591 | 2.3283 | 2 | 2.1867 | 3.0909 | 3.6467 | 2.5267 | 2.7933 | 2.29
| Median |  | 2.525 | 2.8428 | 3.65 | 3.455 | 3.655 | 3.1706 | 2.305 | 2 | 2.195 | 3.0967 | 3.64 | 2.52 | 2.75 | 2.295
| Std Dev. |  | 0.045935 | 0.17134 | 0.17889 | 0.14128 | 0.080166 | 0.040811 | 0.075741 | 0 | 0.065625 | 0.1779 | 0.064704 | 0.072019 | 0.17002 | 0.047329
| Rel.Std.Dev. |  | 1.819% | 6.049% | 4.969% | 4.155% | 2.188% | 1.292% | 3.253% | 0% | 3.001% | 5.756% | 1.774% | 2.85% | 6.087% | 2.067%
| PDM3 |  | -13.71% | -3.212% | 23.02% | 16.19% | 25.19% | 7.954% | -20.44% | -31.66% | -25.28% | 5.622% | 24.62% | -13.66% | -4.545% | -21.74%
| Z-Score (Absolute) |  | 0.41 | 0 | 1.03 | 0.76 | 1.11 | 0.44 | 0.67 | NA | 0.86 | 0.35 | 1.09 | 0.41 | 0.05 | 0.73
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 262b (Indicative Value 2.01 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS |  | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g |  | 2
| 1 | 1 | 1.01: 1.915; 1.02: 2.1112; 1.03: 2.2 |  | 1
| 1 | 2 | 1.01: 1.931; 1.02: 2.0753; 1.03: 1.9 |  | 8
| 1 | 3 | 1.01: 1.937; 1.02: 2.0759; 1.03: 2 |  | 16
| 1 | 4 | 1.01: 1.903; 1.02: 2.0176; 1.03: 1.9 |  | 2.0099
| 1 | 5 | 1.01: 1.967; 1.02: 2.1437; 1.03: 2 |  | 14
| 1 | 6 | 1.01: 1.92; 1.02: 2.1823; 1.03: 2
| Mean |  | 1.01: 1.9288; 1.02: 2.101; 1.03: 2
| Median |  | 1.01: 1.9255; 1.02: 2.0935; 1.03: 2
| Std Dev. |  | 1.01: 0.022203; 1.02: 0.057955; 1.03: 0.10954
| Rel.Std.Dev. |  | 1.01: 1.151%; 1.02: 2.758%; 1.03: 5.477%
| PDM3 |  | 1.01: -4.035%; 1.02: 4.53%; 1.03: -0.4947%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.96; 1.03: 0
| Analytical results for Ge in OREAS 262b (Indicative Value 0.06 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.03; 1.02: 1.09; 1.03: 1.10; 1.05: 1.12; 1.06: 1.15; 1.07: 1.19; 1.09: 1.20; 1.10: 1.21; 1.11: 1.22 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*OES/MS; 1.03: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*MS; 1.09: AR*MS; 1.10: AR*MS; 1.11: AR*MS |  | 3
|  |  | 1.01: 0.2g; 1.02: 0.5g; 1.03: 0.2g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.25g |  | 3
| 1 | 1 | 1.01: < 0.2; 1.02: < 0.1; 1.03: < 0.05; 1.05: < 0.05; 1.06: 0.05; 1.07: 0.05; 1.09: 0.15; 1.10: 0.05; 1.11: < 0.1 |  | 1
| 1 | 2 | 1.01: < 0.2; 1.02: < 0.1; 1.03: < 0.05; 1.05: < 0.05; 1.06: 0.05; 1.07: 0.05; 1.09: 0.12; 1.10: 0.05; 1.11: < 0.1 |  | 3
| 1 | 3 | 1.01: < 0.2; 1.02: < 0.1; 1.03: < 0.05; 1.05: < 0.05; 1.06: 0.05; 1.07: 0.05; 1.09: 0.1; 1.10: < 0.05; 1.11: < 0.1 |  | 16
| 1 | 4 | 1.01: < 0.2; 1.02: < 0.1; 1.03: < 0.05; 1.05: < 0.05; 1.06: 0.06; 1.07: 0.05; 1.09: 0.09; 1.10: 0.05; 1.11: < 0.1 |  | 0.059708
| 1 | 5 | 1.01: < 0.2; 1.02: < 0.1; 1.03: < 0.05; 1.05: < 0.05; 1.06: 0.05; 1.07: 0.05; 1.09: 0.09; 1.10: < 0.05; 1.11: < 0.1 |  | 15
| 1 | 6 | 1.01: < 0.2; 1.02: < 0.1; 1.03: < 0.05; 1.05: < 0.05; 1.06: 0.06; 1.07: 0.05; 1.09: 0.09; 1.10: < 0.05; 1.11: < 0.1
| Mean |  | 1.06: 0.053333; 1.07: 0.05; 1.09: 0.10667; 1.10: 0.05
| Median |  | 1.06: 0.05; 1.07: 0.05; 1.09: 0.095; 1.10: 0.05
| Std Dev. |  | 1.06: 0.005164; 1.07: 7.6012e-18; 1.09: 0.024221; 1.10: 8.4984e-18
| Rel.Std.Dev. |  | 1.06: 9.682%; 1.07: 1.52e-14%; 1.09: 22.71%; 1.10: 1.7e-14%
| PDM3 |  | 1.06: -10.68%; 1.07: -16.26%; 1.09: 78.65%; 1.10: -16.26%
| Z-Score (Absolute) |  | 1.01: 2.7; 1.02: 0; 1.03: 1.35; 1.05: 1.35; 1.06: 0.18; 1.07: 0; 1.09: 3.06; 1.10: 0.67; 1.11: 0
| Analytical results for Hf in OREAS 262b (Certified Value 0.33 ppm) | [20 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.33 | 0.34182 | 0.37 | 0.3 | 0.26 | 0.39 | 0.34 | 0.30457 | 0.29 | 0.31 | 0.35 | 0.34 | 0.17 | [6 empty] | 1
| 1 | 2 | 0.3 | 0.34855 | 0.35 | 0.3 | 0.29 | 0.39 | 0.35 | 0.33095 | 0.31 | 0.33 | 0.34 | 0.33 | 0.18 | [6 empty] | 21
| 1 | 3 | 0.3 | 0.34895 | 0.39 | 0.3 | 0.27 | 0.39 | 0.34 | 0.33144 | 0.28 | 0.31 | 0.31 | 0.32 | 0.18 | [6 empty] | 16
| 1 | 4 | 0.31 | 0.33762 | 0.35 | 0.2 | 0.29 | 0.36 | 0.33 | 0.33183 | 0.31 | 0.33 | 0.32 | 0.33 | 0.18 | [6 empty] | 0.32979
| 1 | 5 | 0.3 | 0.34299 | 0.34 | 0.2 | 0.34 | 0.38 | 0.35 | 0.33168 | 0.28 | 0.33 | 0.36 | 0.32 | 0.18 | [6 empty] | 90
| 1 | 6 | 0.3 | 0.36298 | 0.33 | 0.3 | 0.3 | 0.34 | 0.35 | 0.34184 | 0.25 | 0.34 | 0.37 | 0.32 | 0.2
| Mean |  | 0.30667 | 0.34715 | 0.355 | 0.26667 | 0.29167 | 0.375 | 0.34333 | 0.32872 | 0.28667 | 0.325 | 0.34167 | 0.32667 | 0.18167
| Median |  | 0.3 | 0.34577 | 0.35 | 0.3 | 0.29 | 0.385 | 0.345 | 0.33156 | 0.285 | 0.33 | 0.345 | 0.325 | 0.18
| Std Dev. |  | 0.012111 | 0.0088568 | 0.021679 | 0.05164 | 0.027869 | 0.020736 | 0.008165 | 0.012539 | 0.022509 | 0.012247 | 0.023166 | 0.008165 | 0.0098319
| Rel.Std.Dev. |  | 3.949% | 2.551% | 6.107% | 19.36% | 9.555% | 5.53% | 2.378% | 3.814% | 7.852% | 3.768% | 6.78% | 2.499% | 5.412%
| PDM3 |  | -7.012% | 5.264% | 7.644% | -19.14% | -11.56% | 13.71% | 4.106% | -0.3249% | -13.08% | -1.453% | 3.601% | -0.9474% | -44.91%
| Z-Score (Absolute) |  | 0.7 | 0.65 | 0.91 | NA | 1.2 | 1.58 | 0.52 | 0.03 | 1.37 | 0.09 | 0.47 | 0.03 | 4.86
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 262b (Certified Value 0.056 ppm) | [20 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [5 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | < 0.1 | < 1 | 0.03 | < 2 | 0.06 | 0.05 | 0.042 | 0.06 | < 0.05 | 1 | 0.05 | 0.05 | 0.07 | 0.37 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 1 | 0.05 | < 2 | 0.06 | 0.07 | 0.049 | 0.05 | < 0.05 | 1 | 0.05 | 0.03 | 0.07 | 0.36 | [5 empty] | 3
| 1 | 3 | < 0.1 | < 1 | 0.04 | < 2 | 0.07 | 0.1 | 0.05 | 0.04 | < 0.05 | 1 | 0.05 | 0.04 | 0.07 | 0.39 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 1 | 0.05 | < 2 | 0.06 | < 0.01 | 0.045 | 0.05 | < 0.05 | 1 | 0.06 | 0.05 | 0.06 | 0.38 | [5 empty] | 0.055583
| 1 | 5 | < 0.1 | < 1 | 0.07 | < 2 | 0.07 | 0.08 | 0.048 | 0.05 | < 0.05 | 1 | 0.06 | 0.08 | 0.07 | 0.32 | [5 empty] | 91
| 1 | 6 | < 0.1 | < 1 | 0.03 | < 2 | 0.06 | 0.04 | 0.054 | 0.05 | < 0.05 | 1 | 0.06 | 0.08 | 0.07 | 0.35
| Mean |  | 1.03: 0.045; 1.06: 0.063333; 1.07: 0.068; 1.09: 0.048; 1.10: 0.05; 1.12: 1; 1.13: 0.055; 1.14: 0.055; 1.15: 0.068333; 1.16: 0.36167
| Median |  | 1.03: 0.045; 1.06: 0.06; 1.07: 0.07; 1.09: 0.0485; 1.10: 0.05; 1.12: 1; 1.13: 0.055; 1.14: 0.05; 1.15: 0.07; 1.16: 0.365
| Std Dev. |  | 1.03: 0.015166; 1.06: 0.005164; 1.07: 0.023875; 1.09: 0.0041473; 1.10: 0.0063246; 1.12: 0; 1.13: 0.0054772; 1.14: 0.020736; 1.15: 0.0040825; 1.16: 0.024833
| Rel.Std.Dev. |  | 1.03: 33.7%; 1.06: 8.154%; 1.07: 35.11%; 1.09: 8.64%; 1.10: 12.65%; 1.12: 0%; 1.13: 9.959%; 1.14: 37.7%; 1.15: 5.974%; 1.16: 6.866%
| PDM3 |  | 1.03: -19.04%; 1.06: 13.94%; 1.07: 22.34%; 1.09: -13.64%; 1.10: -10.04%; 1.12: 1699%; 1.13: -1.049%; 1.14: -1.049%; 1.15: 22.94%; 1.16: 550.7%
| Z-Score (Absolute) |  | 0.4 | 36.01 | 0.81 | 76.47 | 0.67 | 0.2 | 0.57 | 0.4 | 2.43 | NA | 0 | 0 | 1.08 | 24.81
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 262b (Indicative Value 0.19 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS |  | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g |  | 2
| 1 | 1 | 1.01: 0.199; 1.02: 0.19364; 1.03: 0.2 |  | 1
| 1 | 2 | 1.01: 0.19; 1.02: 0.18839; 1.03: 0.2 |  | 10
| 1 | 3 | 1.01: 0.193; 1.02: 0.18054; 1.03: 0.2 |  | 16
| 1 | 4 | 1.01: 0.191; 1.02: 0.19111; 1.03: 0.2 |  | 0.19434
| 1 | 5 | 1.01: 0.195; 1.02: 0.18837; 1.03: 0.2 |  | 16
| 1 | 6 | 1.01: 0.192; 1.02: 0.19615; 1.03: 0.2
| Mean |  | 1.01: 0.19333; 1.02: 0.1897; 1.03: 0.2
| Median |  | 1.01: 0.1925; 1.02: 0.18975; 1.03: 0.2
| Std Dev. |  | 1.01: 0.003266; 1.02: 0.0054131; 1.03: 3.0405e-17
| Rel.Std.Dev. |  | 1.01: 1.689%; 1.02: 2.854%; 1.03: 1.52e-14%
| PDM3 |  | 1.01: -0.5203%; 1.02: -2.39%; 1.03: 2.91%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 1.24
| Analytical results for In in OREAS 262b (Certified Value 0.019 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.02 | < 0.05 | 0.02 | < 5 | 0.018 | 0.015 | < 0.05 | 0.1 | 0.017 | 0.019 | 0.019 | 0.022 | [7 empty] | 1
| 1 | 2 | 0.02 | < 0.05 | 0.02 | < 5 | 0.017 | 0.022 | < 0.05 | 0.1 | 0.019 | 0.024 | 0.019 | 0.02 | [7 empty] | 23
| 1 | 3 | 0.02 | < 0.05 | 0.03 | < 5 | 0.018 | 0.017 | < 0.05 | 0.1 | 0.018 | 0.02 | 0.019 | 0.024 | [7 empty] | 16
| 1 | 4 | 0.02 | < 0.05 | 0.02 | < 5 | 0.018 | 0.019 | < 0.05 | 0.1 | 0.016 | 0.023 | 0.019 | 0.02 | [7 empty] | 0.019342
| 1 | 5 | 0.02 | < 0.05 | 0.03 | < 5 | 0.017 | 0.019 | < 0.05 | 0.1 | 0.018 | 0.022 | 0.018 | 0.02 | [7 empty] | 92
| 1 | 6 | 0.02 | < 0.05 | 0.02 | < 5 | 0.02 | 0.017 | < 0.05 | 0.1 | 0.018 | 0.024 | 0.017 | 0.02
| Mean |  | 1.01: 0.02; 1.03: 0.023333; 1.06: 0.018; 1.07: 0.018167; 1.10: 0.1; 1.11: 0.017667; 1.12: 0.022; 1.13: 0.0185; 1.14: 0.021
| Median |  | 1.01: 0.02; 1.03: 0.02; 1.06: 0.018; 1.07: 0.018; 1.10: 0.1; 1.11: 0.018; 1.12: 0.0225; 1.13: 0.019; 1.14: 0.02
| Std Dev. |  | 1.01: 0; 1.03: 0.005164; 1.06: 0.0010954; 1.07: 0.0024014; 1.10: 1.5202e-17; 1.11: 0.0010328; 1.12: 0.0020976; 1.13: 0.00083666; 1.14: 0.0016733
| Rel.Std.Dev. |  | 1.01: 0%; 1.03: 22.13%; 1.06: 6.086%; 1.07: 13.22%; 1.10: 1.52e-14%; 1.11: 5.846%; 1.12: 9.535%; 1.13: 4.522%; 1.14: 7.968%
| PDM3 |  | 1.01: 3.404%; 1.03: 20.64%; 1.06: -6.937%; 1.07: -6.075%; 1.10: 417%; 1.11: -8.66%; 1.12: 13.74%; 1.13: -4.352%; 1.14: 8.574%
| Z-Score (Absolute) |  | 0.24 | 0.95 | 0.56 | 589.98 | 0.71 | 0.67 | 0.95 | NA | 0.79 | 0.24 | 0.59 | 0
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 262b (Certified Value 0.216 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.2221 | 0.21707 | 0.18334 | 0.35 | 0.234 | 0.23 | 0.2 | 0.22 | 0.2 | 0.21 | 0.19 | 0.2475 | 0.29 | 0.21 | 0.19 | 0.2 | 0.221 |  |  | 1
| 1 | 2 | 0.2192 | 0.21724 | 0.18385 | 0.35 | 0.236 | 0.24 | 0.21 | 0.22 | 0.2 | 0.19 | 0.18 | 0.24745 | 0.29 | 0.21 | 0.18 | 0.19 | 0.227 |  |  | 0
| 1 | 3 | 0.2083 | 0.21464 | 0.18207 | 0.33 | 0.24 | 0.24 | 0.2 | 0.22 | 0.21 | 0.2 | 0.18 | 0.24887 | 0.29 | 0.21 | 0.19 | 0.19 | 0.215 |  |  | 16
| 1 | 4 | 0.215 | 0.21462 | 0.1844 | 0.35 | 0.262 | 0.24 | 0.21 | 0.22 | 0.2 | 0.2 | 0.19 | 0.25681 | 0.29 | 0.21 | 0.19 | 0.19 | 0.216 |  |  | 0.21649
| 1 | 5 | 0.2196 | 0.22172 | 0.18522 | 0.34 | 0.244 | 0.24 | 0.21 | 0.22 | 0.2 | 0.2 | 0.18 | 0.25976 | 0.29 | 0.22 | 0.19 | 0.2 | 0.204 |  |  | 93
| 1 | 6 | 0.2134 | 0.2182 | 0.182 | 0.33 | 0.249 | 0.24 | 0.21 | 0.22 | 0.19 | 0.2 | 0.19 | 0.25775 | 0.28 | 0.22 | 0.19 | 0.19 | 0.215
| Mean |  | 0.21627 | 0.21725 | 0.18348 | 0.34167 | 0.24417 | 0.23833 | 0.20667 | 0.22 | 0.2 | 0.2 | 0.185 | 0.25302 | 0.28833 | 0.21333 | 0.18833 | 0.19333 | 0.21633
| Median |  | 0.2171 | 0.21715 | 0.18359 | 0.345 | 0.242 | 0.24 | 0.21 | 0.22 | 0.2 | 0.2 | 0.185 | 0.25284 | 0.29 | 0.21 | 0.19 | 0.19 | 0.2155
| Std Dev. |  | 0.0050405 | 0.0026335 | 0.0012818 | 0.0098319 | 0.010284 | 0.0040825 | 0.005164 | 0 | 0.0063246 | 0.0063246 | 0.0054772 | 0.0056722 | 0.0040825 | 0.005164 | 0.0040825 | 0.005164 | 0.0076333
| Rel.Std.Dev. |  | 2.331% | 1.212% | 0.6986% | 2.878% | 4.212% | 1.713% | 2.499% | 0% | 3.162% | 3.162% | 2.961% | 2.242% | 1.416% | 2.421% | 2.168% | 2.671% | 3.528%
| PDM3 |  | -0.1035% | 0.3496% | -15.25% | 57.82% | 12.78% | 10.09% | -4.538% | 1.621% | -7.617% | -7.617% | -14.55% | 16.88% | 33.19% | -1.458% | -13.01% | -10.7% | -0.07267%
| Z-Score (Absolute) |  | 0 | 0.03 | 1 | 3.83 | 0.85 | 0.67 | 0.29 | 0.11 | 0.5 | 0.5 | 0.96 | 1.12 | 2.2 | 0.09 | 0.85 | 0.7 | 0
| Analytical results for La in OREAS 262b (Certified Value 14.5 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 14.027 | 13.8861 | 15.3 | 16.8 | 15.2 | 14.7167 | 13.8 | 15 | 13.8 | 15.9648 | 11.9 | 14.3 | 16.6 | 13.7 | 14.8 |  |  |  |  | 1
| 1 | 2 | 14.004 | 14.0141 | 15.4 | 15.3 | 14.9 | 14.065 | 13.7 | 14 | 13.6 | 15.7739 | 11.6 | 14.6 | 14.7 | 13.9 | 15 |  |  |  |  | 24
| 1 | 3 | 14.399 | 13.7878 | 15.6 | 16.2 | 14.7 | 14.71 | 14.3 | 13 | 12.9 | 15.779 | 12 | 14.2 | 14.3 | 13.8 | 14.8 |  |  |  |  | 16
| 1 | 4 | 13.936 | 13.7769 | 15.5 | 15.5 | 14.9 | 14.8967 | 13.9 | 14 | 14 | 16.1056 | 11.8 | 14 | 15.8 | 13.7 | 14.5 |  |  |  |  | 14.489
| 1 | 5 | 14.389 | 14.1686 | 15.5 | 16 | 15.3 | 14.3667 | 13.8 | 15 | 13.6 | 16.0229 | 12.3 | 14.8 | 15 | 13.4 | 14.7 |  |  |  |  | 94
| 1 | 6 | 14.078 | 13.8173 | 15.5 | 16.3 | 15 | 14.6167 | 13.7 | 14 | 14 | 16.0923 | 11.8 | 15.1 | 16.6 | 12.7 | 14.8
| Mean |  | 14.1388 | 13.9085 | 15.4667 | 16.0167 | 15 | 14.5619 | 13.8667 | 14.1667 | 13.65 | 15.9564 | 11.9 | 14.5 | 15.5 | 13.5333 | 14.7667
| Median |  | 14.0525 | 13.8517 | 15.5 | 16.1 | 14.95 | 14.6633 | 13.8 | 14 | 13.7 | 15.9938 | 11.85 | 14.45 | 15.4 | 13.7 | 14.8
| Std Dev. |  | 0.20288 | 0.1546 | 0.10328 | 0.54924 | 0.21909 | 0.29866 | 0.22509 | 0.75277 | 0.40866 | 0.14831 | 0.23664 | 0.40988 | 0.98387 | 0.44121 | 0.1633
| Rel.Std.Dev. |  | 1.435% | 1.112% | 0.6678% | 3.429% | 1.461% | 2.051% | 1.623% | 5.314% | 2.994% | 0.9295% | 1.989% | 2.827% | 6.348% | 3.26% | 1.106%
| PDM3 |  | -2.417% | -4.006% | 6.748% | 10.54% | 3.527% | 0.5036% | -4.295% | -2.225% | -5.79% | 10.13% | -17.87% | 0.07608% | 6.978% | -6.596% | 1.917%
| Z-Score (Absolute) |  | 0.34 | 0.54 | 0.82 | 1.3 | 0.41 | 0.03 | 0.58 | NA | 0.77 | 1.24 | 2.3 | 0.03 | 0.85 | 0.87 | 0.21
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 262b (Certified Value 11.3 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 11.6 | 12.7966 | 11 | 13 | 13.8 | 9.5 | 12.1135 | 10.6 | 10 | 7.8993 | 13.7 | 11.6 | 10.2 | 10.7 | 14 |  |  |  |  | 1
| 1 | 2 | 11.4 | 12.5124 | 11 | 13 | 12.5 | 9.5 | 11.76 | 10.6 | 9.8 | 8.0112 | 13.8 | 11.6 | 10 | 10.6 | 14 |  |  |  |  | 0
| 1 | 3 | 11.5 | 12.4647 | 10 | 12.7 | 13.2 | 9.5 | 11.9729 | 11.2 | 10 | 7.4788 | 13.8 | 11.5 | 10.3 | 10.6 | 13 |  |  |  |  | 16
| 1 | 4 | 11.5 | 12.7317 | 12 | 13.1 | 13 | 9.2 | 12.0619 | 11 | 10.2 | 8.0156 | 14.5 | 11.2 | 10.4 | 10.6 | 13 |  |  |  |  | 11.3032
| 1 | 5 | 11.9 | 13.1402 | 11 | 13.2 | 13.7 | 9.4 | 11.9357 | 10.8 | 9.9 | 7.4875 | 14.1 | 11.9 | 10.3 | 10.7 | 13 |  |  |  |  | 95
| 1 | 6 | 11.8 | 12.6164 | 11 | 13.1 | 13.4 | 10.6 | 12.0953 | 10.5 | 10.1 | 7.6329 | 13.4 | 12 | 10.2 | 10.6 | 13
| Mean |  | 11.6167 | 12.7103 | 11 | 13.0167 | 13.2667 | 9.6167 | 11.9899 | 10.7833 | 10 | 7.7542 | 13.8833 | 11.6333 | 10.2333 | 10.6333 | 13.3333
| Median |  | 11.55 | 12.6741 | 11 | 13.05 | 13.3 | 9.5 | 12.0174 | 10.7 | 10 | 7.7661 | 13.8 | 11.6 | 10.25 | 10.6 | 13
| Std Dev. |  | 0.19408 | 0.24534 | 0.63246 | 0.17224 | 0.48028 | 0.49565 | 0.1324 | 0.27142 | 0.14142 | 0.25184 | 0.37639 | 0.28752 | 0.13663 | 0.05164 | 0.5164
| Rel.Std.Dev. |  | 1.671% | 1.93% | 5.75% | 1.323% | 3.62% | 5.154% | 1.104% | 2.517% | 1.414% | 3.248% | 2.711% | 2.472% | 1.335% | 0.4856% | 3.873%
| PDM3 |  | 2.774% | 12.45% | -2.682% | 15.16% | 17.37% | -14.92% | 6.075% | -4.599% | -11.53% | -31.4% | 22.83% | 2.921% | -9.465% | -5.926% | 17.96%
| Z-Score (Absolute) |  | 0 | 0.53 | NA | 0.68 | 0.8 | 0.97 | 0.18 | 0.41 | 0.79 | 1.88 | 1.1 | 0.01 | 0.67 | 0.48 | NA
Results from laboratories 1.06 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 262b (Indicative Value 0.06 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.05: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09; 1.05: 1.22 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS; 1.05: AR*MS |  | 3
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g; 1.05: 0.25g |  | 3
| 1 | 1 | 1.01: 0.056; 1.02: 0.062756; 1.03: < 0.1; 1.05: 0.06 |  | 1
| 1 | 2 | 1.01: 0.054; 1.02: 0.060749; 1.03: < 0.1; 1.05: 0.06 |  | 5
| 1 | 3 | 1.01: 0.055; 1.02: 0.062978; 1.03: < 0.1; 1.05: 0.07 |  | 16
| 1 | 4 | 1.01: 0.055; 1.02: 0.061096; 1.03: < 0.1; 1.05: 0.06 |  | 0.059761
| 1 | 5 | 1.01: 0.053; 1.02: 0.065614; 1.03: < 0.1; 1.05: 0.06 |  | 11
| 1 | 6 | 1.01: 0.055; 1.02: 0.064501; 1.03: < 0.1; 1.05: 0.06
| Mean |  | 1.01: 0.054667; 1.02: 0.062949; 1.05: 0.061667
| Median |  | 1.01: 0.055; 1.02: 0.062867; 1.05: 0.06
| Std Dev. |  | 1.01: 0.0010328; 1.02: 0.0018878; 1.05: 0.0040825
| Rel.Std.Dev. |  | 1.01: 1.889%; 1.02: 2.999%; 1.05: 6.62%
| PDM3 |  | 1.01: -8.524%; 1.02: 5.335%; 1.05: 3.189%
| Z-Score (Absolute) |  | 1.01: 0.57; 1.02: 0.78; 1.03: 1.33; 1.05: 0.57
| Analytical results for Mg in OREAS 262b (Certified Value 0.287 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.29 | 0.38141 | 0.27166 | 0.23473 | 0.57 | 0.31 | 0.39 | 0.27 | 0.29 | 0.25 | 0.29 | 0.23 | 0.32048 | 0.34 | 0.26 | 0.24 | 0.25 | 0.282 |  | 1
| 1 | 2 | 0.29 | 0.38932 | 0.27552 | 0.23477 | 0.58 | 0.32 | 0.36 | 0.26 | 0.26 | 0.25 | 0.26 | 0.23 | 0.31417 | 0.34 | 0.26 | 0.23 | 0.24 | 0.287 |  | 0
| 1 | 3 | 0.27 | 0.39389 | 0.26905 | 0.23419 | 0.55 | 0.33 | 0.38 | 0.26 | 0.29 | 0.26 | 0.28 | 0.23 | 0.31311 | 0.34 | 0.26 | 0.24 | 0.24 | 0.273 |  | 16
| 1 | 4 | 0.28 | 0.38201 | 0.27215 | 0.2363 | 0.6 | 0.32 | 0.36 | 0.26 | 0.29 | 0.25 | 0.28 | 0.23 | 0.33628 | 0.34 | 0.26 | 0.24 | 0.24 | 0.276 |  | 0.28656
| 1 | 5 | 0.28 | 0.38826 | 0.28108 | 0.23971 | 0.55 | 0.32 | 0.4 | 0.27 | 0.29 | 0.25 | 0.28 | 0.23 | 0.32876 | 0.34 | 0.27 | 0.24 | 0.24 | 0.283 |  | 96
| 1 | 6 | 0.28 | 0.38317 | 0.27194 | 0.23454 | 0.57 | 0.31 | 0.37 | 0.27 | 0.29 | 0.26 | 0.28 | 0.23 | 0.32752 | 0.33 | 0.27 | 0.24 | 0.24 | 0.273
| Mean |  | 0.28167 | 0.38634 | 0.27357 | 0.23571 | 0.57 | 0.31833 | 0.37667 | 0.265 | 0.285 | 0.25333 | 0.27833 | 0.23 | 0.32339 | 0.33833 | 0.26333 | 0.23833 | 0.24167 | 0.279
| Median |  | 0.28 | 0.38572 | 0.27205 | 0.23475 | 0.57 | 0.32 | 0.375 | 0.265 | 0.29 | 0.25 | 0.28 | 0.23 | 0.324 | 0.34 | 0.26 | 0.24 | 0.24 | 0.279
| Std Dev. |  | 0.0075277 | 0.004952 | 0.0042169 | 0.0020899 | 0.018974 | 0.0075277 | 0.01633 | 0.0054772 | 0.012247 | 0.005164 | 0.0098319 | 0 | 0.0090689 | 0.0040825 | 0.005164 | 0.0040825 | 0.0040825 | 0.005831
| Rel.Std.Dev. |  | 2.673% | 1.282% | 1.541% | 0.8866% | 3.329% | 2.365% | 4.335% | 2.067% | 4.297% | 2.038% | 3.532% | 0% | 2.804% | 1.207% | 1.961% | 1.713% | 1.689% | 2.09%
| PDM3 |  | -1.707% | 34.82% | -4.534% | -17.75% | 98.91% | 11.09% | 31.44% | -7.523% | -0.5441% | -11.59% | -2.871% | -19.74% | 12.85% | 18.07% | -8.105% | -16.83% | -15.67% | -2.638%
| Z-Score (Absolute) |  | 0.05 | 1.89 | 0.09 | 0.76 | 5.12 | 0.7 | 1.72 | 0.24 | 0.11 | 0.45 | 0.01 | 0.86 | 0.79 | 1.05 | 0.27 | 0.71 | 0.65 | 0.01
| Analytical results for Mn in OREAS 262b (Certified Value 0.014 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0149 | 0.014939 | 0.014372 | 0.012871 | 0.02 | 0.0147 | 0.0149 | 0.0145 | 0.0169 | 0.0145 | 0.0156 | 0.0134 | 0.014249 | 0.0143 | 0.0145 | 0.0135 | 0.0138 | 0.0138 |  | 1
| 1 | 2 | 0.0148 | 0.015032 | 0.014496 | 0.01341 | 0.02 | 0.0149 | 0.0126 | 0.0143 | 0.0165 | 0.0147 | 0.0145 | 0.0129 | 0.014372 | 0.0143 | 0.0143 | 0.0133 | 0.0135 | 0.0141 |  | 26
| 1 | 3 | 0.0139 | 0.015418 | 0.01427 | 0.013396 | 0.02 | 0.0152 | 0.0131 | 0.0143 | 0.0169 | 0.0149 | 0.0147 | 0.0133 | 0.014346 | 0.0143 | 0.0144 | 0.0135 | 0.0137 | 0.0134 |  | 16
| 1 | 4 | 0.0147 | 0.014985 | 0.014291 | 0.013485 | 0.02 | 0.0151 | 0.0126 | 0.0144 | 0.0169 | 0.015 | 0.0149 | 0.0135 | 0.015203 | 0.0146 | 0.0144 | 0.0136 | 0.0136 | 0.0134 |  | 0.014201
| 1 | 5 | 0.0146 | 0.015277 | 0.014567 | 0.013316 | 0.02 | 0.0149 | 0.0139 | 0.0147 | 0.0167 | 0.0148 | 0.0152 | 0.013 | 0.014895 | 0.0145 | 0.0149 | 0.0133 | 0.0138 | 0.0128 |  | 97
| 1 | 6 | 0.0143 | 0.014745 | 0.014511 | 0.013353 | 0.02 | 0.0148 | 0.0134 | 0.0142 | 0.0168 | 0.0144 | 0.015 | 0.0133 | 0.015198 | 0.0141 | 0.0153 | 0.0134 | 0.0136 | 0.0133
| Mean |  | 0.014533 | 0.015066 | 0.014418 | 0.013305 | 0.02 | 0.014933 | 0.013417 | 0.0144 | 0.016783 | 0.014717 | 0.014983 | 0.013233 | 0.014711 | 0.01435 | 0.014633 | 0.013433 | 0.013667 | 0.013467
| Median |  | 0.01465 | 0.015008 | 0.014434 | 0.013374 | 0.02 | 0.0149 | 0.01325 | 0.01435 | 0.01685 | 0.01475 | 0.01495 | 0.0133 | 0.014634 | 0.0143 | 0.01445 | 0.01345 | 0.01365 | 0.0134
| Std Dev. |  | 0.00037238 | 0.00024298 | 0.00012418 | 0.00022008 | 0 | 0.00018619 | 0.00087958 | 0.00017889 | 0.00016021 | 0.00023166 | 0.00038687 | 0.00023381 | 0.0004414 | 0.00017607 | 0.00038816 | 0.00012111 | 0.00012111 | 0.00044572
| Rel.Std.Dev. |  | 2.562% | 1.613% | 0.8613% | 1.654% | 0% | 1.247% | 6.556% | 1.242% | 0.9546% | 1.574% | 2.582% | 1.767% | 3.001% | 1.227% | 2.653% | 0.9015% | 0.8861% | 3.31%
| PDM3 |  | 2.339% | 6.088% | 1.524% | -6.31% | 40.83% | 5.155% | -5.524% | 1.4% | 18.18% | 3.63% | 5.507% | -6.815% | 3.587% | 1.048% | 3.043% | -5.407% | -3.764% | -5.172%
| Z-Score (Absolute) |  | 0.07 | 0.72 | 0.07 | 1.44 | 6.78 | 0.56 | 1.3 | 0.09 | 2.83 | 0.3 | 0.62 | 1.53 | 0.29 | 0.15 | 0.19 | 1.28 | 0.99 | 1.24
| Analytical results for Mo in OREAS 262b (Certified Value 1.39 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 1.5 | 1.6548 | < 10 | 1.4 | 1.47 | 1.36 | < 2 | 1.33 | 1.5 | 1.521 | 1.22 | 1.2303 | 1.41 | 1.18 | 1.8 | 1.21 | 1.72 |  |  | 1
| 1 | 2 | 1.6 | 1.6334 | < 10 | 1.4 | 1.27 | 1.32 | < 2 | 1.32 | 1.4 | 1.538 | 1.63 | 1.3828 | 1.55 | 1.23 | 1.36 | 1.44 | 1.94 |  |  | 27
| 1 | 3 | 1.5 | 1.2852 | < 10 | 1.5 | 1.42 | 1.27 | < 2 | 1.39 | 1 | 1.473 | 1.18 | 1.3836 | 1.56 | 1.21 | 1.43 | 1.26 | 1.85 |  |  | 16
| 1 | 4 | 1.5 | 1.5949 | < 10 | 1.3 | 1.32 | 1.29 | < 2 | 1.38 | 0.8 | 1.498 | 1.26 | 1.3451 | 1.4 | 1.24 | 1.44 | 1.23 | 1.57 |  |  | 1.3885
| 1 | 5 | 1.5 | 1.3892 | < 10 | 1.4 | 1.39 | 1.35 | < 2 | 1.29 | 1.3 | 1.665 | 1.18 | 1.31 | 1.4 | 1.24 | 1.37 | 1.18 | 1.45 |  |  | 98
| 1 | 6 | 1.6 | 1.6465 | < 10 | 1.3 | 1.34 | 1.41 | < 2 | 1.42 | 0.9 | 1.646 | 1.3 | 1.4011 | 1.44 | 1.38 | 1.44 | 1.22 | 1.59
| Mean |  | 1.5333 | 1.534 |  | 1.3833 | 1.3683 | 1.3333 |  | 1.355 | 1.15 | 1.5568 | 1.295 | 1.3422 | 1.46 | 1.2467 | 1.4733 | 1.2567 | 1.6867
| Median |  | 1.5 | 1.6142 |  | 1.4 | 1.365 | 1.335 |  | 1.355 | 1.15 | 1.5295 | 1.24 | 1.364 | 1.425 | 1.235 | 1.435 | 1.225 | 1.655
| Std Dev. |  | 0.05164 | 0.15731 |  | 0.075277 | 0.072503 | 0.050859 |  | 0.049295 | 0.2881 | 0.079723 | 0.17062 | 0.063888 | 0.0751 | 0.069186 | 0.16391 | 0.093524 | 0.1849
| Rel.Std.Dev. |  | 3.368% | 10.25% |  | 5.442% | 5.299% | 3.814% |  | 3.638% | 25.05% | 5.121% | 13.18% | 4.76% | 5.144% | 5.55% | 11.13% | 7.442% | 10.96%
| PDM3 |  | 10.43% | 10.48% |  | -0.3702% | -1.45% | -3.971% |  | -2.411% | -17.18% | 12.13% | -6.732% | -3.334% | 5.151% | -10.21% | 6.112% | -9.493% | 21.48%
| Z-Score (Absolute) |  | NA | 1.07 | 22.65 | NA | 0.04 | 0.18 | 2.25 | 0.04 | NA | 1.21 | 0.41 | 0.12 | 0.61 | 0.72 | 0.7 | 0.65 | 2.02
Results from laboratories 1.01, 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 262b (Indicative Value 0.01 wt.%) | [20 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | < 0.01 | < 0.01 | < 0.01 | 0.96 | 0.01 | 0.016 | 0.01 | 0.01 | 0.01 | 0.008 | < 0.01 | 0.013284 | 0.01 | < 0.01 | 0.02 | 0.01 | < 0.005 |  |  | 1
| 1 | 2 | < 0.01 | < 0.01 | < 0.01 | 0.94 | 0.01 | 0.015 | 0.01 | 0.01 | 0.01 | 0.007 | < 0.01 | 0.012536 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.005 |  |  | 6
| 1 | 3 | < 0.01 | 0.010072 | < 0.01 | 0.93 | 0.01 | 0.016 | 0.01 | 0.01 | 0.01 | 0.007 | < 0.01 | 0.013926 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.005 |  |  | 16
| 1 | 4 | < 0.01 | < 0.01 | < 0.01 | 0.98 | 0.01 | 0.017 | 0.01 | 0.01 | 0.01 | 0.007 | < 0.01 | 0.014392 | 0.01 | < 0.01 | 0.02 | 0.01 | < 0.005 |  |  | 0.01001
| 1 | 5 | < 0.01 | 0.010385 | < 0.01 | 0.9 | 0.01 | 0.017 | 0.01 | 0.01 | 0.01 | 0.008 | < 0.01 | 0.012593 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.005 |  |  | 12
| 1 | 6 | < 0.01 | < 0.01 | < 0.01 | 0.97 | 0.01 | 0.017 | 0.01 | 0.01 | < 0.01 | 0.008 | < 0.01 | 0.013706 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.005
| Mean |  |  | 0.010229 |  | 0.94667 | 0.01 | 0.016333 | 0.01 | 0.01 | 0.01 | 0.0075 |  | 0.013406 | 0.01 |  | 0.013333 | 0.01
| Median |  |  | 0.010229 |  | 0.95 | 0.01 | 0.0165 | 0.01 | 0.01 | 0.01 | 0.0075 |  | 0.013495 | 0.01 |  | 0.01 | 0.01
| Std Dev. |  |  | 0.00022186 |  | 0.029439 | 0 | 0.0008165 | 0 | 0 | 0 | 0.00054772 |  | 0.00074383 | 0 |  | 0.005164 | 0
| Rel.Std.Dev. |  |  | 2.169% |  | 3.11% | 0% | 4.999% | 0% | 0% | 0% | 7.303% |  | 5.548% | 0% |  | 38.73% | 0%
| PDM3 |  |  | 2.181% |  | 9357% | -0.1023% | 63.17% | -0.1023% | -0.1023% | -0.1023% | -25.08% |  | 33.92% | -0.1023% |  | 33.2% | -0.1023%
| Z-Score (Absolute) |  | 1.01 | 0.66 | 1.01 | 189.48 | 0 | 1.28 | 0 | 0 | 0.17 | 0.51 | 1.01 | 0.69 | 0 | 1.01 | 0.67 | 0 | 1.52
| Analytical results for Nb in OREAS 262b (Indicative Value 0.17 ppm) | [20 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.25 | 0.093135 | 22 | 0.3 | 0.1 | 0.27 | < 5 | 0.13 | 0.11 | < 0.2 | 0.2 | 0.12 | 0.21 | 0.11 | 0.15 |  |  |  |  | 1
| 1 | 2 | 0.23 | 0.10106 | 17 | 0.3 | < 0.1 | 0.27 | < 5 | 0.14 | 0.11 | 0.20148 | 0.2 | 0.13 | 0.16 | 0.1 | 0.18 |  |  |  |  | 7
| 1 | 3 | 0.25 | 0.095534 | 20 | 0.3 | 0.1 | 0.24 | < 5 | 0.13 | 0.1 | < 0.2 | 0.2 | 0.13 | 0.14 | 0.11 | 0.17 |  |  |  |  | 16
| 1 | 4 | 0.24 | 0.097748 | 17 | 0.3 | < 0.1 | 0.26 | < 5 | 0.14 | 0.11 | 0.20292 | 0.2 | 0.14 | 0.16 | 0.1 | 0.16 |  |  |  |  | 0.16931
| 1 | 5 | 0.25 | 0.099977 | 11 | 0.3 | < 0.1 | 0.28 | < 5 | 0.13 | 0.1 | < 0.2 | 0.2 | 0.13 | 0.15 | 0.09 | 0.16 |  |  |  |  | 13
| 1 | 6 | 0.24 | 0.10187 | 15 | 0.3 | < 0.1 | 0.28 | < 5 | 0.13 | 0.11 | < 0.2 | 0.2 | 0.13 | 0.16 | 0.09 | 0.18
| Mean |  | 0.24333 | 0.09822 | 17 | 0.3 | 0.1 | 0.26667 |  | 0.13333 | 0.10667 | 0.2022 | 0.2 | 0.13 | 0.16333 | 0.1 | 0.16667
| Median |  | 0.245 | 0.098863 | 17 | 0.3 | 0.1 | 0.27 |  | 0.13 | 0.11 | 0.2022 | 0.2 | 0.13 | 0.16 | 0.1 | 0.165
| Std Dev. |  | 0.008165 | 0.0033982 | 3.8471 | 0 | 0 | 0.015055 |  | 0.005164 | 0.005164 | 0.0010207 | 3.0405e-17 | 0.0063246 | 0.024221 | 0.0089443 | 0.012111
| Rel.Std.Dev. |  | 3.355% | 3.46% | 22.63% | 0% | 0% | 5.646% |  | 3.873% | 4.841% | 0.5048% | 1.52e-14% | 4.865% | 14.83% | 8.944% | 7.266%
| PDM3 |  | 43.72% | -41.99% | 9941% | 77.19% | -40.94% | 57.5% |  | -21.25% | -37% | 19.42% | 18.12% | -23.22% | -3.532% | -40.94% | -1.564%
| Z-Score (Absolute) |  | 0.85 | 0.69 | 179.26 | 1.46 | 1.03 | 1.1 | 24.88 | 0.32 | 0.6 | 0.31 | 0.39 | 0.35 | 0 | 0.67 | 0.04
| Analytical results for Nd in OREAS 262b (Indicative Value 14.2 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.09 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*OES/MS |  | 1
|  |  | 1.01: 25g; 1.02: 0.5g |  | 1
| 1 | 1 | 1.01: 13.487; 1.02: 16.3 |  | 1
| 1 | 2 | 1.01: 13.373; 1.02: 14.4 |  | 8
| 1 | 3 | 1.01: 13.486; 1.02: 15.2 |  | 16
| 1 | 4 | 1.01: 13.455; 1.02: 13.9 |  | 14.232
| 1 | 5 | 1.01: 13.6; 1.02: 15.2 |  | 14
| 1 | 6 | 1.01: 13.383; 1.02: 15
| Mean |  | 1.01: 13.464; 1.02: 15
| Median |  | 1.01: 13.4705; 1.02: 15.1
| Std Dev. |  | 1.01: 0.082984; 1.02: 0.81731
| Rel.Std.Dev. |  | 1.01: 0.6163%; 1.02: 5.449%
| PDM3 |  | 1.01: -5.396%; 1.02: 5.396%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ni in OREAS 262b (Certified Value 16 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 17 | 16.8156 | 16.2384 | 11.853 | 17 | 16 | 20.7 | 17.4 | 16.2653 | 15.6 | 15 | 15.5 | 17.8029 | 18.1 | 14.7 | 16.2 | 17.6 | 15.6 |  | 1
| 1 | 2 | 16.6 | 18.2304 | 15.4121 | 16.104 | 16 | 16 | 20.7 | 17.3 | 15.881 | 15.5 | 14.3 | 14.2 | 17.6352 | 17.9 | 15.6 | 15.3 | 15 | 15 |  | 30
| 1 | 3 | 16.3 | 17.008 | 14.9289 | 14.239 | 14 | 16 | 19.2 | 17.1 | 16.1467 | 16.8 | 15.1 | 13.6 | 17.9481 | 17.9 | 15.6 | 15.4 | 16 | 15.8 |  | 16
| 1 | 4 | 16.7 | 16.9452 | 14.9822 | 13.233 | 139 | 16 | 20.1 | 17.4 | 16.1643 | 15.6 | 15.2 | 15.2 | 17.7968 | 17.3 | 15.8 | 17.3 | 15.9 | 14.6 |  | 15.9961
| 1 | 5 | 16.6 | 17.442 | 17.7558 | 14.228 | 16 | 16 | 20.3 | 17.7 | 15.9463 | 15.1 | 15.1 | 14.5 | 17.4607 | 17.8 | 17.1 | 16.8 | 14.8 | 15.4 |  | 99
| 1 | 6 | 16.4 | 16.5456 | 15.6214 | 13.172 | 15 | 16 | 20.8 | 17.1 | 16.071 | 15.3 | 16.1 | 14.3 | 17.6665 | 17.5 | 16.1 | 16.8 | 14.7 | 15.9
| Mean |  | 16.6 | 17.1645 | 15.8232 | 13.8048 | 36.1667 | 16 | 20.3 | 17.3333 | 16.0791 | 15.65 | 15.1333 | 14.55 | 17.7184 | 17.75 | 15.8167 | 16.3 | 15.6667 | 15.3833
| Median |  | 16.6 | 16.9766 | 15.5168 | 13.7305 | 16 | 16 | 20.5 | 17.35 | 16.1088 | 15.55 | 15.1 | 14.4 | 17.7316 | 17.85 | 15.7 | 16.5 | 15.45 | 15.5
| Std Dev. |  | 0.24495 | 0.59841 | 1.0606 | 1.4276 | 50.3882 | 0 | 0.60332 | 0.22509 | 0.14384 | 0.59582 | 0.57504 | 0.69498 | 0.16856 | 0.29496 | 0.78337 | 0.81486 | 1.0985 | 0.49967
| Rel.Std.Dev. |  | 1.476% | 3.486% | 6.703% | 10.34% | 139.3% | 0% | 2.972% | 1.299% | 0.8946% | 3.807% | 3.8% | 4.777% | 0.9513% | 1.662% | 4.953% | 4.999% | 7.012% | 3.248%
| PDM3 |  | 3.775% | 7.304% | -1.081% | -13.7% | 126.1% | 0.02431% | 26.91% | 8.36% | 0.5189% | -2.164% | -5.394% | -9.04% | 10.77% | 10.96% | -1.122% | 1.9% | -2.06% | -3.831%
| Z-Score (Absolute) |  | 0.48 | 0.97 | 0.19 | 1.93 | 17.37 | 0.03 | 3.68 | 1.12 | 0.03 | 0.34 | 0.78 | 1.29 | 1.45 | 1.48 | 0.19 | 0.22 | 0.32 | 0.57
| Analytical results for P in OREAS 262b (Certified Value 0.019 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0199 | 0.018984 | 0.015225 | 0.02 | 0.0205 | 0.021 | 0.0191 | 0.02 | 0.019 | 0.02 | 0.018 | 0.01921 | 0.02 | 0.019 | 0.019 | 0.019 | 0.018 |  |  | 1
| 1 | 2 | 0.0197 | 0.019663 | 0.01634 | 0.02 | 0.0202 | 0.021 | 0.0185 | 0.0193 | 0.019 | 0.019 | 0.018 | 0.019144 | 0.02 | 0.019 | 0.018 | 0.019 | 0.019 |  |  | 0
| 1 | 3 | 0.0184 | 0.018944 | 0.015713 | 0.02 | 0.0204 | 0.021 | 0.0188 | 0.0197 | 0.02 | 0.02 | 0.019 | 0.020339 | 0.02 | 0.019 | 0.019 | 0.019 | 0.019 |  |  | 16
| 1 | 4 | 0.0197 | 0.018993 | 0.015676 | 0.02 | 0.0206 | 0.021 | 0.0189 | 0.0198 | 0.02 | 0.02 | 0.019 | 0.020481 | 0.02 | 0.019 | 0.019 | 0.019 | 0.019 |  |  | 0.019492
| 1 | 5 | 0.0195 | 0.02003 | 0.015554 | 0.02 | 0.0201 | 0.021 | 0.0189 | 0.0197 | 0.02 | 0.02 | 0.018 | 0.02144 | 0.02 | 0.02 | 0.019 | 0.019 | 0.017 |  |  | 100
| 1 | 6 | 0.019 | 0.019204 | 0.016158 | 0.02 | 0.0205 | 0.021 | 0.0186 | 0.0198 | 0.019 | 0.02 | 0.018 | 0.021177 | 0.02 | 0.02 | 0.018 | 0.019 | 0.018
| Mean |  | 0.019367 | 0.019303 | 0.015778 | 0.02 | 0.020383 | 0.021 | 0.0188 | 0.019717 | 0.0195 | 0.019833 | 0.018333 | 0.020298 | 0.02 | 0.019333 | 0.018667 | 0.019 | 0.018333
| Median |  | 0.0196 | 0.019098 | 0.015695 | 0.02 | 0.02045 | 0.021 | 0.01885 | 0.01975 | 0.0195 | 0.02 | 0.018 | 0.02041 | 0.02 | 0.019 | 0.019 | 0.019 | 0.0185
| Std Dev. |  | 0.00056451 | 0.00044591 | 0.00040754 | 0 | 0.00019408 | 0 | 0.00021909 | 0.00023166 | 0.00054772 | 0.00040825 | 0.0005164 | 0.00096182 | 0 | 0.0005164 | 0.0005164 | 0 | 0.0008165
| Rel.Std.Dev. |  | 2.915% | 2.31% | 2.583% | 0% | 0.9521% | 0% | 1.165% | 1.175% | 2.809% | 2.058% | 2.817% | 4.738% | 0% | 2.671% | 2.766% | 0% | 4.454%
| PDM3 |  | -0.6417% | -0.9683% | -19.05% | 2.608% | 4.574% | 7.738% | -3.549% | 1.154% | 0.04232% | 1.752% | -5.943% | 4.138% | 2.608% | -0.8127% | -4.233% | -2.523% | -5.943%
| Z-Score (Absolute) |  | 0 | 0.07 | 3.82 | 0.67 | 1.08 | 1.74 | 0.6 | 0.37 | 0.14 | 0.5 | 1.1 | 0.99 | 0.67 | 0.04 | 0.75 | 0.39 | 1.1
| Analytical results for Pb in OREAS 262b (Certified Value 16.6 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 16.6 | 16.192 | < 100 | 28 | 16 | 17.7 | 16.3 | 17.4319 | 16.4 | 15.6 | 11.5891 | 15.4 | 16.8545 | 17.8 | 16.7 | 17.8 | 16.1 | 18.4 |  | 1
| 1 | 2 | 15.9 | 16.0496 | < 100 | 49 | 15.7 | 17.1 | 16.2 | 17.2702 | 16 | 15.2 | 11.3284 | 15.4 | 17.2869 | 17.7 | 17 | 17.2 | 16.5 | 17 |  | 31
| 1 | 3 | 16.3 | 15.8332 | < 100 | 24 | 16.4 | 16.7 | 16.6 | 17.5654 | 16.2 | 13.4 | 11.7905 | 14.8 | 17.2932 | 17.5 | 16.6 | 16.5 | 16.7 | 18.6 |  | 16
| 1 | 4 | 16.2 | 15.748 | < 100 | 32 | 16.1 | 16.7 | 15.9 | 17.6561 | 15.7 | 13.8 | 11.145 | 15.9 | 17.8522 | 17.9 | 16.4 | 17 | 16.2 | 18.9 |  | 16.6464
| 1 | 5 | 15.8 | 16.8769 | < 100 | 32 | 16.3 | 16.8 | 16.5 | 18.3516 | 16.3 | 15.9 | 11.7926 | 15.3 | 18.2222 | 18 | 17.4 | 16.8 | 16.4 | 18.9 |  | 101
| 1 | 6 | 16.1 | 16.0141 | < 100 | 23 | 16.1 | 17.8 | 16.7 | 17.3527 | 16.4 | 14.6 | 11.5359 | 15.7 | 18.314 | 17.3 | 17.4 | 17.6 | 16.1 | 17.3
| Mean |  | 16.15 | 16.119 |  | 31.3333 | 16.1 | 17.1333 | 16.3667 | 17.6046 | 16.1667 | 14.75 | 11.5303 | 15.4167 | 17.6371 | 17.7 | 16.9167 | 17.15 | 16.3333 | 18.1833
| Median |  | 16.15 | 16.0318 |  | 30 | 16.1 | 16.95 | 16.4 | 17.4987 | 16.25 | 14.9 | 11.5625 | 15.4 | 17.5727 | 17.75 | 16.85 | 17.1 | 16.3 | 18.5
| Std Dev. |  | 0.2881 | 0.40371 |  | 9.4587 | 0.24495 | 0.50067 | 0.29439 | 0.39173 | 0.27325 | 0.9995 | 0.2566 | 0.37639 | 0.58311 | 0.26077 | 0.42151 | 0.48888 | 0.24221 | 0.82805
| Rel.Std.Dev. |  | 1.784% | 2.505% |  | 30.19% | 1.521% | 2.922% | 1.799% | 2.225% | 1.69% | 6.776% | 2.225% | 2.441% | 3.306% | 1.473% | 2.492% | 2.851% | 1.483% | 4.554%
| PDM3 |  | -2.982% | -3.169% |  | 88.23% | -3.283% | 2.925% | -1.681% | 5.756% | -2.882% | -11.39% | -30.73% | -7.388% | 5.952% | 6.329% | 1.623% | 3.025% | -1.881% | 9.233%
| Z-Score (Absolute) |  | 0.19 | 0.21 | 28.95 | NA | 0.23 | 0.66 | 0 | 1.07 | 0.17 | 1.39 | 4.16 | 0.82 | 1.09 | 1.15 | 0.47 | 0.67 | 0.03 | 1.56
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 262b (Indicative Value < 10 ppb) | [20 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.16 |  | ppb
|  |  | 1.01: AR*MS; 1.03: AR*OES/MS |  | 1
|  |  | 1.01: 25g; 1.03: 01g |  | 1
| 1 | 1 | 1.01: < 10; 1.03: < 10 |  | 1
| 1 | 2 | 1.01: < 10; 1.03: < 10 |  | 1
| 1 | 3 | 1.01: < 10; 1.03: < 10 |  | 16
| 1 | 4 | 1.01: < 10; 1.03: < 10 |  | < 10
| 1 | 5 | 1.01: < 10; 1.03: < 10 |  | 15
| 1 | 6 | 1.01: < 10; 1.03: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Pr in OREAS 262b (Indicative Value 3.81 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.05: 1.09 |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.05: AR*OES/MS |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.05: 0.5g |  | 2
| 1 | 1 | 1.01: 3.708; 1.03: 3.7202; 1.05: 4.4 |  | 1
| 1 | 2 | 1.01: 3.481; 1.03: 3.6123; 1.05: 4 |  | 10
| 1 | 3 | 1.01: 3.626; 1.03: 3.6081; 1.05: 4.2 |  | 16
| 1 | 4 | 1.01: 3.488; 1.03: 3.6606; 1.05: 3.9 |  | 3.8074
| 1 | 5 | 1.01: 3.522; 1.03: 3.7513; 1.05: 4.2 |  | 16
| 1 | 6 | 1.01: 3.533; 1.03: 3.8222; 1.05: 4.3
| Mean |  | 1.01: 3.5597; 1.03: 3.6958; 1.05: 4.1667
| Median |  | 1.01: 3.5275; 1.03: 3.6904; 1.05: 4.2
| Std Dev. |  | 1.01: 0.089283; 1.03: 0.084328; 1.05: 0.18619
| Rel.Std.Dev. |  | 1.01: 2.508%; 1.03: 2.282%; 1.05: 4.469%
| PDM3 |  | 1.01: -6.506%; 1.03: -2.931%; 1.05: 9.437%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0; 1.05: 2.33
| Analytical results for Pt in OREAS 262b (Indicative Value < 5 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.16 |  | ppb
|  |  | 1.01: AR*MS; 1.03: AR*OES/MS |  | 2
|  |  | 1.01: 25g; 1.03: 01g |  | 2
| 1 | 1 | 1.01: < 5; 1.03: < 5 |  | 1
| 1 | 2 | 1.01: < 5; 1.03: < 5 |  | 1
| 1 | 3 | 1.01: < 5; 1.03: < 5 |  | 16
| 1 | 4 | 1.01: < 5; 1.03: < 5 |  | < 5
| 1 | 5 | 1.01: < 5; 1.03: < 5 |  | 11
| 1 | 6 | 1.01: < 5; 1.03: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Rb in OREAS 262b (Certified Value 17 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 16.96 | 16.7472 | 17.3 | 17.9 | 18.6 | 15.4 | 15.8 | 18.0385 | 22.7 | 16.3 | 17.3 | 15.8 | 18.65 | [6 empty] | 1
| 1 | 2 | 16.37 | 16.9444 | 17.5 | 16.5 | 18.7 | 16.2 | 15.2 | 18.088 | 22.8 | 16.8 | 15.8 | 15.7 | 17.56 | [6 empty] | 33
| 1 | 3 | 16.8 | 16.9503 | 17.6 | 17.7 | 18.8 | 16.7 | 14.9 | 17.998 | 22.8 | 16.4 | 15.9 | 15.5 | 18.05 | [6 empty] | 16
| 1 | 4 | 16.22 | 16.592 | 17.3 | 17 | 18.7 | 15.4 | 16.3 | 18.436 | 23.4 | 16.2 | 16.7 | 15.7 | 18.27 | [6 empty] | 16.9634
| 1 | 5 | 16.84 | 17.2952 | 17.6 | 18.4 | 19 | 15.6 | 15.4 | 18.432 | 23.4 | 17.1 | 15.8 | 15.5 | 17.75 | [6 empty] | 102
| 1 | 6 | 16.78 | 17.1286 | 17.6 | 17.9 | 18.4 | 16.4 | 15.8 | 18.2135 | 22.6 | 17.1 | 16.4 | 15 | 17.65
| Mean |  | 16.6617 | 16.9429 | 17.4833 | 17.5667 | 18.7 | 15.95 | 15.5667 | 18.201 | 22.95 | 16.65 | 16.3167 | 15.5333 | 17.9883
| Median |  | 16.79 | 16.9474 | 17.55 | 17.8 | 18.7 | 15.9 | 15.6 | 18.1507 | 22.8 | 16.6 | 16.15 | 15.6 | 17.9
| Std Dev. |  | 0.29465 | 0.25303 | 0.1472 | 0.69186 | 0.2 | 0.55767 | 0.50067 | 0.1945 | 0.35637 | 0.40373 | 0.6047 | 0.28752 | 0.41849
| Rel.Std.Dev. |  | 1.768% | 1.493% | 0.8419% | 3.938% | 1.07% | 3.496% | 3.216% | 1.069% | 1.553% | 2.425% | 3.706% | 1.851% | 2.326%
| PDM3 |  | -1.779% | -0.1205% | 3.065% | 3.556% | 10.24% | -5.974% | -8.234% | 7.296% | 35.29% | -1.847% | -3.812% | -8.43% | 6.042%
| Z-Score (Absolute) |  | 0.19 | 0 | 0.37 | 0.42 | 1.19 | 0.67 | 0.93 | 0.85 | 4.08 | 0.2 | 0.43 | 0.96 | 0.71
| Analytical results for Re in OREAS 262b (Certified Value < 0.001 ppm) | [20 empty] | Certified
Re | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.07; 1.05: 1.09; 1.06: 1.10; 1.07: 1.12; 1.09: 1.15; 1.10: 1.16; 1.11: 1.19; 1.12: 1.20; 1.13: 1.21 |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*OES/MS; 1.05: AR*OES/MS; 1.06: AR*MS; 1.07: AR*MS; 1.09: AR*MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS |  | 3
|  |  | 1.01: 25g; 1.03: 01g; 1.05: 0.5g; 1.06: 0.2g; 1.07: 0.5g; 1.09: 0.5g; 1.10: 01g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g |  | 3
| 1 | 1 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: < 0.001; 1.10: < 0.05; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001 |  | 1
| 1 | 2 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.002; 1.06: < 0.002; 1.07: < 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: < 0.001; 1.12: < 0.001; 1.13: 0.001 |  | 34
| 1 | 3 | 1.01: < 0.001; 1.03: < 0.05; 1.05: < 0.001; 1.06: < 0.002; 1.07: < 0.001; 1.09: < 0.001; 1.10: < 0.05; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001 |  | 16
| 1 | 4 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: < 0.001; 1.10: < 0.05; 1.11: < 0.001; 1.12: < 0.001; 1.13: 0.001 |  | < 0.001
| 1 | 5 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.002; 1.06: < 0.002; 1.07: < 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: < 0.001; 1.12: 0.001; 1.13: < 0.001 |  | 103
| 1 | 6 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: < 0.001; 1.10: < 0.05; 1.11: < 0.001; 1.12: 0.001; 1.13: 0.001
| Mean |  | 1.05: 0.0014; 1.07: 0.001; 1.09: 0.001; 1.12: 0.001; 1.13: 0.001
| Median |  | 1.05: 0.001; 1.07: 0.001; 1.09: 0.001; 1.12: 0.001; 1.13: 0.001
| Std Dev. |  | 1.05: 0.00054772; 1.07: 0; 1.09: 0; 1.12: 0; 1.13: 0
| Rel.Std.Dev. |  | 1.05: 39.12%; 1.07: 0%; 1.09: 0%; 1.12: 0%; 1.13: 0%
PDM3
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 65.41; 1.05: 1.35; 1.06: 0.67; 1.07: 0; 1.09: 0.22; 1.10: 65.41; 1.11: 0.67; 1.12: 0.22; 1.13: 0
| Analytical results for S in OREAS 262b (Certified Value 0.031 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0335 | 0.033002 |  | 0.6 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.030772 | 0.06 | 0.032838 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 |  | 1
| 1 | 2 | 0.0331 | 0.032648 |  | 0.6 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.03097 | 0.07 | 0.046184 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 |  | 19
| 1 | 3 | 0.0326 | 0.033252 |  | 0.58 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.030869 | 0.07 | 0.033867 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 |  | 16
| 1 | 4 | 0.0331 | 0.032525 |  | 0.63 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.030671 | 0.03 | 0.037444 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 |  | 0.030629
| 1 | 5 | 0.0328 | 0.034436 |  | 0.58 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.030869 | 0.03 | 0.04535 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 |  | 104
| 1 | 6 | 0.0322 | 0.033537 |  | 0.59 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.030667 | 0.05 | 0.040321 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03
| Mean |  | 0.032883 | 0.033233 |  | 0.59667 | 0.03 |  | 0.03 | 0.03 | 0.03 |  | 0.030803 | 0.051667 | 0.039334 | 0.03 | 0.03 | 0.03 | 0.04 | 0.031667
| Median |  | 0.03295 | 0.033127 |  | 0.595 | 0.03 |  | 0.03 | 0.03 | 0.03 |  | 0.03082 | 0.055 | 0.038882 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03
| Std Dev. |  | 0.0004535 | 0.00069804 |  | 0.018619 | 0 |  | 0 | 0 | 0 |  | 0.00012107 | 0.018348 | 0.0056526 | 0 | 0 | 0 | 0 | 0.0040825
| Rel.Std.Dev. |  | 1.379% | 2.1% |  | 3.121% | 0% |  | 0% | 0% | 0% |  | 0.393% | 35.51% | 14.37% | 0% | 0% | 0% | 0% | 12.89%
| PDM3 |  | 7.36% | 8.503% |  | 1848% | -2.054% |  | -2.054% | -2.054% | -2.054% |  | 0.5676% | 68.69% | 28.42% | -2.054% | -2.054% | -2.054% | 30.59% | 3.388%
| Z-Score (Absolute) |  | 1.17 | 1.31 | NA | 229.27 | 0 | 11.94 | 0 | 0 | 0 | 2.02 | 0.32 | 8.77 | 3.78 | 0 | 0 | 0 | 4.05 | 0.67
| Analytical results for Sb in OREAS 262b (Certified Value 3.55 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 4.16 | 3.0885 | < 10 | 4.5 | 3.77 | 4.09 | 3.66 | < 5 | 3.49 | 3.7 | 2.991 | 2.88 | 3.4812 | 4.54 | 3.05 | 4.09 | 3.1 | 2.97 |  | 1
| 1 | 2 | 4.1 | 3.0698 | < 10 | 4.7 | 3.58 | 3.44 | 3.66 | < 5 | 3.53 | 3.8 | 3.193 | 2.82 | 3.3755 | 4.53 | 3.19 | 3.82 | 3.01 | 3 |  | 35
| 1 | 3 | 4.06 | 3.2281 | < 10 | 3.3 | 3.7 | 3.81 | 3.6 | < 5 | 3.65 | 4 | 3.204 | 2.75 | 3.2644 | 4.55 | 3.12 | 3.77 | 3.14 | 3.05 |  | 16
| 1 | 4 | 4.02 | 3.0103 | < 10 | 3.7 | 3.63 | 3.58 | 3.69 | < 5 | 3.46 | 4 | 3.073 | 3 | 3.4641 | 4.66 | 3.23 | 3.91 | 3.15 | 2.94 |  | 3.5503
| 1 | 5 | 4.1 | 3.338 | < 10 | 3.8 | 3.57 | 3.81 | 3.74 | < 5 | 3.59 | 3.6 | 3.234 | 2.83 | 3.2737 | 4.62 | 3.16 | 3.73 | 2.97 | 3.01 |  | 105
| 1 | 6 | 4.14 | 3.2103 | < 10 | 6 | 3.66 | 3.75 | 3.68 | < 5 | 3.69 | 4.2 | 3.488 | 2.94 | 3.4207 | 4.45 | 3.16 | 3.69 | 2.99 | 2.89
| Mean |  | 4.0967 | 3.1575 |  | 4.3333 | 3.6517 | 3.7467 | 3.6717 |  | 3.5683 | 3.8833 | 3.1972 | 2.87 | 3.3799 | 4.5583 | 3.1517 | 3.835 | 3.06 | 2.9767
| Median |  | 4.1 | 3.1494 |  | 4.15 | 3.645 | 3.78 | 3.67 |  | 3.56 | 3.9 | 3.1985 | 2.855 | 3.3981 | 4.545 | 3.16 | 3.795 | 3.055 | 2.985
| Std Dev. |  | 0.051251 | 0.12204 |  | 0.96885 | 0.075741 | 0.22259 | 0.045789 |  | 0.090866 | 0.22286 | 0.16966 | 0.089889 | 0.093426 | 0.073598 | 0.061779 | 0.14639 | 0.079498 | 0.056451
| Rel.Std.Dev. |  | 1.251% | 3.865% |  | 22.36% | 2.074% | 5.941% | 1.247% |  | 2.546% | 5.739% | 5.307% | 3.132% | 2.764% | 1.615% | 1.96% | 3.817% | 2.598% | 1.896%
| PDM3 |  | 15.39% | -11.06% |  | 22.06% | 2.856% | 5.531% | 3.419% |  | 0.5083% | 9.381% | -9.946% | -19.16% | -4.799% | 28.39% | -11.23% | 8.019% | -13.81% | -16.16%
| Z-Score (Absolute) |  | 0.72 | 0.67 | 2.06 | 1.07 | 0.06 | 0.2 | 0.09 | 1.64 | 0.06 | 0.4 | 0.61 | 1.1 | 0.34 | 1.4 | 0.68 | 0.33 | 0.81 | 0.94
| Analytical results for Sc in OREAS 262b (Certified Value 1.55 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 1.7 | 1.5445 | 2 | 1.7 | 1.7 | < 5 | 1.5 | 1.4 | 1.4 | 1.4558 | 1.9 | 1.6 | 1.6 | 1.6 | 1.4 |  |  |  |  | 1
| 1 | 2 | 1.7 | 1.5187 | 2 | 1.7 | 1.7 | < 5 | 1.5 | 1.4 | 1.4 | 1.4306 | 1.9 | 1.7 | 1.5 | 1.6 | 1.5 |  |  |  |  | 36
| 1 | 3 | 1.6 | 1.5175 | 2 | 1.6 | 1.6 | < 5 | 1.6 | 1.5 | 1.3 | 1.4127 | 1.9 | 1.6 | 1.4 | 1.6 | 1.6 |  |  |  |  | 16
| 1 | 4 | 1.6 | 1.5236 | 2 | 1.6 | 1.7 | < 5 | 1.4 | 1.5 | 1.4 | 1.6028 | 1.9 | 1.6 | 1.6 | 1.6 | 1.5 |  |  |  |  | 1.5522
| 1 | 5 | 1.7 | 1.57 | 2 | 1.8 | 1.7 | < 5 | 1.5 | 1.5 | 1.4 | 1.6219 | 1.9 | 1.7 | 1.5 | 1.6 | 1.4 |  |  |  |  | 106
| 1 | 6 | 1.6 | 1.5376 | 2 | 1.8 | 1.6 | < 5 | 1.3 | 1.5 | 1.4 | 1.5234 | 1.9 | 1.7 | 1.6 | 1.5 | 1.5
| Mean |  | 1.65 | 1.5353 | 2 | 1.7 | 1.6667 |  | 1.4667 | 1.4667 | 1.3833 | 1.5079 | 1.9 | 1.65 | 1.5333 | 1.5833 | 1.4833
| Median |  | 1.65 | 1.5306 | 2 | 1.7 | 1.7 |  | 1.5 | 1.5 | 1.4 | 1.4896 | 1.9 | 1.65 | 1.55 | 1.6 | 1.5
| Std Dev. |  | 0.054772 | 0.020135 | 0 | 0.089443 | 0.05164 |  | 0.10328 | 0.05164 | 0.040825 | 0.089435 | 2.4324e-16 | 0.054772 | 0.08165 | 0.040825 | 0.075277
| Rel.Std.Dev. |  | 3.32% | 1.311% | 0% | 5.261% | 3.098% |  | 7.042% | 3.521% | 2.951% | 5.931% | 1.28e-14% | 3.32% | 5.325% | 2.578% | 5.075%
| PDM3 |  | 6.3% | -1.088% | 28.85% | 9.521% | 7.374% |  | -5.511% | -5.511% | -10.88% | -2.857% | 22.41% | 6.3% | -1.216% | 2.005% | -4.437%
| Z-Score (Absolute) |  | 0.67 | 0.18 | NA | 1.03 | 0.79 | 6.92 | 0.68 | 0.68 | 1.29 | 0.38 | 2.51 | 0.67 | 0.19 | 0.18 | 0.56
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 262b (Indicative Value 0.25 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | < 1.5 | < 1 | 0.5 | 0.3 | < 5 | 0.3 | < 0.5 | 0.177 | 0.2 | < 1 | 0.8 | < 0.2 | < 0.2 | < 0.2 |  |  |  |  | 1
| 1 | 2 | < 1 | < 1.5 | < 1 | 0.5 | 0.3 | < 5 | 0.2 | < 0.5 | 0.175 | 0.3 | < 1 | 0.6 | < 0.2 | < 0.2 | 0.2 |  |  |  |  | 37
| 1 | 3 | < 1 | < 1.5 | < 1 | 0.4 | 0.2 | < 5 | 0.2 | < 0.5 | 0.18 | 0.4 | < 1 | 0.8 | 0.2 | < 0.2 | 0.2 |  |  |  |  | 16
| 1 | 4 | < 1 | < 1.5 | < 1 | 0.5 | 0.3 | < 5 | 0.2 | < 0.5 | 0.178 | 0.2 | < 1 | 0.8 | < 0.2 | < 0.2 | 0.2 |  |  |  |  | 0.2475
| 1 | 5 | < 1 | < 1.5 | < 1 | 0.6 | 0.3 | < 5 | 0.5 | < 0.5 | 0.18 | 0.2 | < 1 | 0.6 | < 0.2 | < 0.2 | 0.2 |  |  |  |  | 12
| 1 | 6 | < 1 | < 1.5 | < 1 | 0.4 | 0.3 | < 5 | 0.2 | < 0.5 | 0.185 | 0.2 | < 1 | 0.8 | < 0.2 | 0.4 | < 0.2
| Mean |  |  |  |  | 0.48333 | 0.28333 |  | 0.26667 |  | 0.17917 | 0.25 |  | 0.73333 | 0.2 | 0.4 | 0.2
| Median |  |  |  |  | 0.5 | 0.3 |  | 0.2 |  | 0.179 | 0.2 |  | 0.8 | 0.2 | 0.4 | 0.2
| Std Dev. |  | 1.09: 0.075277; 1.10: 0.040825; 1.12: 0.12111; 1.14: 0.0034303; 1.15: 0.083666; 1.17: 0.10328; 1.21: 0
| Rel.Std.Dev. |  | 1.09: 15.57%; 1.10: 14.41%; 1.12: 45.41%; 1.14: 1.915%; 1.15: 33.47%; 1.17: 14.08%; 1.21: 0%
| PDM3 |  |  |  |  | 95.29% | 14.48% |  | 7.744% |  | -27.61% | 1.01% |  | 196.3% | -19.19% | 61.62% | -19.19%
| Z-Score (Absolute) |  | 0.88 | 1.89 | 0.88 | 0.81 | 0 | 8.97 | 0.07 | 0.13 | 0.42 | 0.13 | 0.88 | 1.82 | 0.67 | 0.54 | 0.47
| Analytical results for Sm in OREAS 262b (Indicative Value 2.62 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g |  |  |  |  | 2
| 1 | 1 | 1.01: 2.606; 1.03: 2.6613; 1.07: 2.5 |  |  |  |  | 1
| 1 | 2 | 1.01: 2.612; 1.03: 2.637; 1.07: 2.4 |  |  |  |  | 7
| 1 | 3 | 1.01: 2.644; 1.03: 2.7014; 1.07: 2.2 |  |  |  |  | 16
| 1 | 4 | 1.01: 2.566; 1.03: 2.5468; 1.07: 2.6 |  |  |  |  | 2.6196
| 1 | 5 | 1.01: 2.58; 1.03: 2.7486; 1.07: 2.9 |  |  |  |  | 13
| 1 | 6 | 1.01: 2.582; 1.03: 2.8678; 1.07: 2.8
| Mean |  | 1.01: 2.5983; 1.03: 2.6938; 1.07: 2.5667
| Median |  | 1.01: 2.594; 1.03: 2.6813; 1.07: 2.55
| Std Dev. |  | 1.01: 0.028211; 1.03: 0.10887; 1.07: 0.2582
| Rel.Std.Dev. |  | 1.01: 1.086%; 1.03: 4.041%; 1.07: 10.06%
| PDM3 |  | 1.01: -0.8119%; 1.03: 2.833%; 1.07: -2.021%
| Z-Score (Absolute) |  | 1.01: 0; 1.03: 2.03; 1.07: 0.67
| Analytical results for Sn in OREAS 262b (Certified Value 1.66 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 1.61 | 1.5788 | 7 | 2 | 1.99 | 1.8 | < 10 | 1.4 | 1.4 | 1.7367 | 2 | 1.5 | 1.8 | 1.5 | 1.5 |  |  |  |  | 1
| 1 | 2 | 1.62 | 1.6413 | 5 | 1.9 | 1.81 | 1.8 | < 10 | 1.5 | 1.3 | 1.6586 | 2.1 | 1.5 | 1.6 | 1.5 | 1.5 |  |  |  |  | 22
| 1 | 3 | 1.64 | 1.6428 | < 1 | 2 | 1.97 | 1.7 | < 10 | 1.5 | 1.3 | 1.5642 | 2 | 1.5 | 1.6 | 1.5 | 1.6 |  |  |  |  | 16
| 1 | 4 | 1.62 | 1.6208 | < 1 | 2 | 1.87 | 1.8 | < 10 | 1.5 | 1.4 | 1.8039 | 2.1 | 1.4 | 1.7 | 1.5 | 1.5 |  |  |  |  | 1.6649
| 1 | 5 | 1.6 | 1.6947 | 3 | 1.9 | 2.01 | 1.8 | < 10 | 1.5 | 1.3 | 1.6049 | 2.1 | 1.5 | 1.7 | 1.4 | 1.5 |  |  |  |  | 107
| 1 | 6 | 1.6 | 1.664 | 1 | 2 | 1.98 | 1.8 | < 10 | 1.5 | 1.4 | 1.5292 | 2 | 1.5 | 1.8 | 1.4 | 1.5
| Mean |  | 1.615 | 1.6404 | 4 | 1.9667 | 1.9383 | 1.7833 |  | 1.4833 | 1.35 | 1.6496 | 2.05 | 1.4833 | 1.7 | 1.4667 | 1.5167
| Median |  | 1.615 | 1.6421 | 4 | 2 | 1.975 | 1.8 |  | 1.5 | 1.35 | 1.6318 | 2.05 | 1.5 | 1.7 | 1.5 | 1.5
| Std Dev. |  | 0.015166 | 0.039233 | 2.582 | 0.05164 | 0.079603 | 0.040825 |  | 0.040825 | 0.054772 | 0.10514 | 0.054772 | 0.040825 | 0.089443 | 0.05164 | 0.040825
| Rel.Std.Dev. |  | 0.9391% | 2.392% | 64.55% | 2.626% | 4.107% | 2.289% |  | 2.752% | 4.057% | 6.374% | 2.672% | 2.752% | 5.261% | 3.521% | 2.692%
| PDM3 |  | -2.995% | -1.47% | 140.3% | 18.13% | 16.43% | 7.115% |  | -10.9% | -18.91% | -0.9181% | 23.13% | -10.9% | 2.11% | -11.91% | -8.902%
| Z-Score (Absolute) |  | 0.13 | 0.02 | NA | 1.34 | 1.22 | 0.58 | 13.99 | 0.67 | 1.23 | 0.02 | 1.69 | 0.67 | 0.23 | 0.74 | 0.54
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 262b (Certified Value 12 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 11.2 | 12.4677 | 186 | 13 | 15.1 | 12.9 | 15.28 | 11.4 | 12 | 11.8 | 11.8036 | 12.3 | 11 | 12.4 | 11.1 | 11.5 |  |  |  | 1
| 1 | 2 | 11.31 | 12.5676 | 184 | 12.4 | 13.4 | 12.5 | 14.73 | 11.5 | 11 | 11.5 | 12.0933 | 12.6 | 11.4 | 11.6 | 11.4 | 11.7 |  |  |  | 23
| 1 | 3 | 11.3 | 12.8454 | 182 | 12.8 | 14.8 | 12.7 | 15.1267 | 11.9 | 12 | 11 | 11.843 | 12.6 | 11.2 | 11.2 | 11.4 | 11.1 |  |  |  | 16
| 1 | 4 | 11.07 | 12.2321 | 191 | 13.3 | 12.9 | 12.7 | 15.32 | 11.6 | 12 | 12.3 | 12.6117 | 13.1 | 11 | 12.2 | 11.4 | 11.1 |  |  |  | 12.0272
| 1 | 5 | 11.28 | 12.7846 | 180 | 12.5 | 14.1 | 12.9 | 15.2667 | 11.5 | 11.8 | 11.6 | 12.3036 | 12.7 | 11.6 | 11.5 | 11.2 | 10.6 |  |  |  | 108
| 1 | 6 | 11.34 | 12.5832 | 183 | 13.1 | 13.3 | 12.5 | 15.43 | 11.6 | 12 | 11.8 | 12.341 | 12.3 | 11.6 | 12.4 | 10.9 | 11.1
| Mean |  | 11.25 | 12.5801 | 184.3333 | 12.85 | 13.9333 | 12.7 | 15.1922 | 11.5833 | 11.8 | 11.6667 | 12.166 | 12.6 | 11.3 | 11.8833 | 11.2333 | 11.1833
| Median |  | 11.29 | 12.5754 | 183.5 | 12.9 | 13.75 | 12.7 | 15.2733 | 11.55 | 12 | 11.7 | 12.1984 | 12.6 | 11.3 | 11.9 | 11.3 | 11.1
| Std Dev. |  | 0.1 | 0.22185 | 3.8297 | 0.35071 | 0.88242 | 0.17889 | 0.2466 | 0.17224 | 0.4 | 0.4274 | 0.31278 | 0.29665 | 0.27568 | 0.51543 | 0.20656 | 0.38166
| Rel.Std.Dev. |  | 0.8889% | 1.764% | 2.078% | 2.729% | 6.333% | 1.409% | 1.623% | 1.487% | 3.39% | 3.663% | 2.571% | 2.354% | 2.44% | 4.337% | 1.839% | 3.413%
| PDM3 |  | -6.462% | 4.597% | 1433% | 6.841% | 15.85% | 5.594% | 26.32% | -3.691% | -1.889% | -2.998% | 1.154% | 4.762% | -6.046% | -1.196% | -6.601% | -7.016%
| Z-Score (Absolute) |  | 0.75 | 0.54 | 165.98 | 0.8 | 1.84 | 0.65 | 3.05 | 0.43 | 0.22 | 0.34 | 0.14 | 0.55 | 0.7 | 0.14 | 0.76 | 0.81
| Analytical results for Ta in OREAS 262b (Certified Value < 0.01 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | 0.06 | [6 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | 0.06 | [6 empty] | 24
| 1 | 3 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | 0.05 | [6 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | 0.05 | [6 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | 0.05 | [6 empty] | 109
| 1 | 6 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | 0.05
| Mean |  | 1.14: 0.3; 1.19: 0.053333
| Median |  | 1.14: 0.3; 1.19: 0.05
| Std Dev. |  | 1.14: 0; 1.19: 0.005164
| Rel.Std.Dev. |  | 1.14: 0%; 1.19: 9.682%
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.19
| Analytical results for Tb in OREAS 262b (Certified Value 0.26 ppm) | [20 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09; 1.09: 1.17; 1.10: 1.22 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g; 1.09: 0.2g; 1.10: 0.25g |  |  |  |  | 2
| 1 | 1 | 1.01: 0.236; 1.03: 0.26091; 1.07: 0.3; 1.09: 0.3; 1.10: 0.27 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.226; 1.03: 0.24214; 1.07: 0.2; 1.09: 0.3; 1.10: 0.27 |  |  |  |  | 25
| 1 | 3 | 1.01: 0.234; 1.03: 0.25148; 1.07: 0.2; 1.09: 0.3; 1.10: 0.26 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.231; 1.03: 0.24967; 1.07: 0.2; 1.09: 0.3; 1.10: 0.27 |  |  |  |  | 0.25911
| 1 | 5 | 1.01: 0.228; 1.03: 0.25787; 1.07: 0.3; 1.09: 0.3; 1.10: 0.24 |  |  |  |  | 110
| 1 | 6 | 1.01: 0.23; 1.03: 0.25612; 1.07: 0.3; 1.09: 0.3; 1.10: 0.26
| Mean |  | 1.01: 0.23083; 1.03: 0.25303; 1.07: 0.25; 1.09: 0.3; 1.10: 0.26167
| Median |  | 1.01: 0.2305; 1.03: 0.2538; 1.07: 0.25; 1.09: 0.3; 1.10: 0.265
| Std Dev. |  | 1.01: 0.0037103; 1.03: 0.0067437; 1.07: 0.054772; 1.09: 0; 1.10: 0.01169
| Rel.Std.Dev. |  | 1.01: 1.607%; 1.03: 2.665%; 1.07: 21.91%; 1.09: 0%; 1.10: 4.468%
| PDM3 |  | 1.01: -10.91%; 1.03: -2.344%; 1.07: -3.515%; 1.09: 15.78%; 1.10: 0.9881%
| Z-Score (Absolute) |  | 1.01: 1.73; 1.03: 0; 1.07: 0.24; 1.09: 3.67; 1.10: 0.67
| Analytical results for Te in OREAS 262b (Certified Value 0.09 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 3
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.1 | < 0.1 | 0.08 | 0.09 | 0.1 | < 5 | 0.08 | < 0.2 | 0.07 | 0.06 | 0.16198 | 0.14 | 0.08 | 0.08 | 0.07 |  |  |  |  | 1
| 1 | 2 | 0.1 | < 0.1 | < 0.05 | 0.11 | 0.11 | < 5 | 0.09 | < 0.2 | 0.07 | 0.07 | 0.21633 | 0.14 | 0.09 | 0.08 | 0.07 |  |  |  |  | 26
| 1 | 3 | 0.1 | 0.1058 | 0.07 | 0.1 | 0.1 | < 5 | 0.1 | < 0.2 | 0.074 | 0.06 | 0.16198 | 0.08 | 0.09 | 0.06 | 0.09 |  |  |  |  | 16
| 1 | 4 | 0.1 | 0.11098 | 0.09 | 0.09 | 0.09 | < 5 | 0.09 | < 0.2 | 0.069 | 0.09 | 0.19211 | 0.11 | 0.08 | 0.08 | 0.08 |  |  |  |  | 0.089796
| 1 | 5 | 0.1 | 0.127 | 0.05 | 0.12 | 0.1 | < 5 | 0.08 | < 0.2 | 0.074 | 0.07 | 0.20172 | 0.1 | 0.09 | 0.09 | 0.09 |  |  |  |  | 111
| 1 | 6 | 0.1 | 0.10378 | 0.09 | 0.1 | 0.09 | < 5 | 0.08 | < 0.2 | 0.071 | 0.09 | 0.22772 | 0.11 | 0.08 | 0.09 | 0.08
| Mean |  | 0.1 | 0.11189 | 0.076 | 0.10167 | 0.098333 |  | 0.086667 |  | 0.071333 | 0.073333 | 0.19364 | 0.11333 | 0.085 | 0.08 | 0.08
| Median |  | 0.1 | 0.10839 | 0.08 | 0.1 | 0.1 |  | 0.085 |  | 0.0705 | 0.07 | 0.19691 | 0.11 | 0.085 | 0.08 | 0.08
| Std Dev. |  | 1.5202e-17 | 0.010519 | 0.016733 | 0.01169 | 0.0075277 |  | 0.008165 |  | 0.0021602 | 0.013663 | 0.027381 | 0.023381 | 0.0054772 | 0.010954 | 0.0089443
| Rel.Std.Dev. |  | 1.52e-14% | 9.401% | 22.02% | 11.5% | 7.655% |  | 9.421% |  | 3.028% | 18.63% | 14.14% | 20.63% | 6.444% | 13.69% | 11.18%
| PDM3 |  | 11.36% | 24.6% | -15.36% | 13.22% | 9.507% |  | -3.485% |  | -20.56% | -18.33% | 115.6% | 26.21% | -5.341% | -10.91% | -10.91%
| Z-Score (Absolute) |  | 0.52 | 0 | 1.42 | 0.62 | 0.42 | 144.26 | 0.28 | 0.52 | 1.19 | 1.07 | 6.13 | 1.32 | 0.37 | 0.67 | 0.67
| Analytical results for Th in OREAS 262b (Certified Value 8.89 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 8.9 | 8.3106 | < 100 | 7.96 | 10 | 6.1 | 9.2 | 8.8 | 8.5 | 9.7916 | 6.68 | 8.5 | 9.8 | 8.3 | 9.41 |  |  |  |  | 1
| 1 | 2 | 8.58 | 8.5819 | < 100 | 7.6 | 9.1 | 6.1 | 8.8 | 8.7 | 8.8 | 9.9249 | 6.02 | 8.7 | 9.3 | 8.6 | 9.31 |  |  |  |  | 27
| 1 | 3 | 8.99 | 8.7108 | < 100 | 8.04 | 9.5 | 6 | 9.2 | 8.5 | 8.6 | 9.6314 | 6.42 | 8.4 | 9 | 8.4 | 9.43 |  |  |  |  | 16
| 1 | 4 | 8.63 | 8.527 | < 100 | 7.73 | 9 | 6.1 | 8.9 | 8.7 | 8.8 | 9.7929 | 6.91 | 8.3 | 9.4 | 8.4 | 9.39 |  |  |  |  | 8.8889
| 1 | 5 | 8.45 | 8.2786 | < 100 | 7.94 | 9.8 | 6.2 | 9.2 | 9.1 | 8.7 | 10.0278 | 6.67 | 8.8 | 9.1 | 8.3 | 9.15 |  |  |  |  | 112
| 1 | 6 | 8.7 | 8.6311 | < 100 | 7.8 | 9.8 | 6.1 | 8.8 | 9.3 | 8.7 | 10.021 | 6.03 | 8.8 | 9.8 | 8.2 | 9.16
| Mean |  | 8.7083 | 8.5067 |  | 7.845 | 9.5333 | 6.1 | 9.0167 | 8.85 | 8.6833 | 9.8649 | 6.455 | 8.5833 | 9.4 | 8.3667 | 9.3083
| Median |  | 8.665 | 8.5544 |  | 7.87 | 9.65 | 6.1 | 9.05 | 8.75 | 8.7 | 9.8589 | 6.545 | 8.6 | 9.35 | 8.35 | 9.35
| Std Dev. |  | 0.20272 | 0.17529 |  | 0.16465 | 0.40825 | 0.063246 | 0.20412 | 0.29496 | 0.1169 | 0.15466 | 0.36741 | 0.2137 | 0.34059 | 0.13663 | 0.12561
| Rel.Std.Dev. |  | 2.328% | 2.061% |  | 2.099% | 4.282% | 1.037% | 2.264% | 3.333% | 1.346% | 1.568% | 5.692% | 2.49% | 3.623% | 1.633% | 1.349%
| PDM3 |  | -2.031% | -4.3% |  | -11.74% | 7.25% | -31.37% | 1.438% | -0.4374% | -2.312% | 10.98% | -27.38% | -3.437% | 5.75% | -5.875% | 4.719%
| Z-Score (Absolute) |  | 0 | 0.23 | 46.41 | 0.97 | 0.93 | 2.93 | 0.35 | 0.16 | 0.03 | 1.3 | 2.53 | 0.14 | 0.78 | 0.38 | 0.67
| Analytical results for Ti in OREAS 262b (Certified Value 0.017 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0165 | 0.018886 | 0.013869 | < 0.1 | 0.0229 | 0.02 | 0.019 | 0.02 | 0.017 | 0.017 | 0.016 | 0.025454 | 0.03 | 0.018 | 0.016 | 0.016 | 0.013 |  |  | 1
| 1 | 2 | 0.0159 | 0.01938 | 0.015172 | < 0.1 | 0.0232 | 0.02 | 0.019 | 0.02 | 0.017 | 0.016 | 0.015 | 0.027027 | 0.03 | 0.018 | 0.015 | 0.016 | 0.014 |  |  | 28
| 1 | 3 | 0.0155 | 0.01903 | 0.014413 | < 0.1 | 0.0237 | 0.02 | 0.019 | 0.02 | 0.018 | 0.014 | 0.015 | 0.025352 | 0.03 | 0.018 | 0.016 | 0.016 | 0.013 |  |  | 16
| 1 | 4 | 0.0155 | 0.019318 | 0.014775 | < 0.1 | 0.0238 | 0.02 | 0.02 | 0.02 | 0.017 | 0.014 | 0.015 | 0.02732 | 0.03 | 0.018 | 0.016 | 0.016 | 0.013 |  |  | 0.017409
| 1 | 5 | 0.0161 | 0.019475 | 0.014996 | < 0.1 | 0.0242 | 0.02 | 0.02 | 0.02 | 0.017 | 0.017 | 0.015 | 0.027732 | 0.03 | 0.019 | 0.016 | 0.016 | 0.012 |  |  | 113
| 1 | 6 | 0.0155 | 0.019551 | 0.014255 | < 0.1 | 0.0244 | 0.02 | 0.021 | 0.02 | 0.016 | 0.014 | 0.016 | 0.027924 | 0.03 | 0.019 | 0.016 | 0.015 | 0.013
| Mean |  | 0.015833 | 0.019273 | 0.01458 |  | 0.0237 | 0.02 | 0.019667 | 0.02 | 0.017 | 0.015333 | 0.015333 | 0.026802 | 0.03 | 0.018333 | 0.015833 | 0.015833 | 0.013
| Median |  | 0.0157 | 0.019349 | 0.014594 |  | 0.02375 | 0.02 | 0.0195 | 0.02 | 0.017 | 0.015 | 0.015 | 0.027173 | 0.03 | 0.018 | 0.016 | 0.016 | 0.013
| Std Dev. |  | 0.00041312 | 0.00026095 | 0.00048987 |  | 0.00057271 | 0 | 0.0008165 | 0 | 0.00063246 | 0.0015055 | 0.0005164 | 0.0011279 | 0 | 0.0005164 | 0.00040825 | 0.00040825 | 0.00063246
| Rel.Std.Dev. |  | 2.609% | 1.354% | 3.36% |  | 2.417% | 0% | 4.152% | 0% | 3.72% | 9.819% | 3.368% | 4.208% | 0% | 2.817% | 2.578% | 2.578% | 4.865%
| PDM3 |  | -9.049% | 10.71% | -16.25% |  | 36.14% | 14.89% | 12.97% | 14.89% | -2.347% | -11.92% | -11.92% | 53.96% | 72.33% | 5.312% | -9.049% | -9.049% | -25.32%
| Z-Score (Absolute) |  | 0.67 | 0.25 | 1.01 | 8.54 | 1.45 | 0.45 | 0.36 | 0.45 | 0.36 | 0.81 | 0.81 | 2.28 | 3.15 | 0 | 0.67 | 0.67 | 1.44
| Analytical results for Tl in OREAS 262b (Certified Value 0.28 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.29 | 0.28338 | 0.29 | 0.34 | 0.28 | < 10 | 0.27 | 0.3 | 0.24 | 0.25 | 0.28016 | 0.36 | 0.26 | 0.31 | 0.27 |  |  |  |  | 1
| 1 | 2 | 0.27 | 0.27751 | 0.3 | 0.31 | 0.28 | < 10 | 0.27 | 0.3 | 0.261 | 0.24 | 0.29335 | 0.37 | 0.27 | 0.3 | 0.26 |  |  |  |  | 29
| 1 | 3 | 0.28 | 0.27228 | 0.3 | 0.32 | 0.28 | < 10 | 0.28 | 0.2 | 0.266 | 0.24 | 0.28927 | 0.38 | 0.25 | 0.28 | 0.26 |  |  |  |  | 16
| 1 | 4 | 0.28 | 0.28408 | 0.29 | 0.31 | 0.27 | < 10 | 0.26 | 0.2 | 0.254 | 0.25 | 0.30354 | 0.36 | 0.25 | 0.29 | 0.25 |  |  |  |  | 0.27935
| 1 | 5 | 0.29 | 0.28576 | 0.29 | 0.34 | 0.28 | < 10 | 0.27 | 0.3 | 0.276 | 0.24 | 0.29282 | 0.36 | 0.27 | 0.29 | 0.26 |  |  |  |  | 114
| 1 | 6 | 0.29 | 0.31146 | 0.29 | 0.33 | 0.3 | < 10 | 0.26 | 0.2 | 0.291 | 0.25 | 0.30231 | 0.35 | 0.27 | 0.3 | 0.26
| Mean |  | 0.28333 | 0.28574 | 0.29333 | 0.325 | 0.28167 |  | 0.26833 | 0.25 | 0.26467 | 0.245 | 0.29357 | 0.36333 | 0.26167 | 0.295 | 0.26
| Median |  | 0.285 | 0.28373 | 0.29 | 0.325 | 0.28 |  | 0.27 | 0.25 | 0.2635 | 0.245 | 0.29308 | 0.36 | 0.265 | 0.295 | 0.26
| Std Dev. |  | 0.008165 | 0.013556 | 0.005164 | 0.013784 | 0.0098319 |  | 0.0075277 | 0.054772 | 0.01766 | 0.0054772 | 0.0086582 | 0.010328 | 0.0098319 | 0.010488 | 0.0063246
| Rel.Std.Dev. |  | 2.882% | 4.744% | 1.76% | 4.241% | 3.491% |  | 2.805% | 21.91% | 6.672% | 2.236% | 2.949% | 2.843% | 3.757% | 3.555% | 2.433%
| PDM3 |  | 1.427% | 2.29% | 5.006% | 16.34% | 0.8301% |  | -3.943% | -10.51% | -5.256% | -12.3% | 5.093% | 30.06% | -6.329% | 5.603% | -6.926%
| Z-Score (Absolute) |  | 0.05 | 0.05 | 0.33 | 1.51 | 0.11 | 176.27 | 0.61 | NA | 0.74 | 1.48 | 0.34 | 2.95 | 0.85 | 0.39 | 0.92
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 262b (Indicative Value 0.067 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS |  |  |  |  | 3
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g |  |  |  |  | 3
| 1 | 1 | 1.01: 0.07; 1.03: 0.067434; 1.07: < 0.1 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.066; 1.03: 0.06166; 1.07: < 0.1 |  |  |  |  | 8
| 1 | 3 | 1.01: 0.072; 1.03: 0.067122; 1.07: < 0.1 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.07; 1.03: 0.064097; 1.07: < 0.1 |  |  |  |  | 0.067115
| 1 | 5 | 1.01: 0.068; 1.03: 0.066766; 1.07: < 0.1 |  |  |  |  | 14
| 1 | 6 | 1.01: 0.068; 1.03: 0.064299; 1.07: < 0.1
| Mean |  | 1.01: 0.069; 1.03: 0.06523
| Median |  | 1.01: 0.069; 1.03: 0.065533
| Std Dev. |  | 1.01: 0.0020976; 1.03: 0.0022669
| Rel.Std.Dev. |  | 1.01: 3.04%; 1.03: 3.475%
| PDM3 |  | 1.01: 2.809%; 1.03: -2.809%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0; 1.07: 2.72
| Analytical results for U in OREAS 262b (Certified Value 1.13 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 1.18 | 1.023 | 1.13 | 1.2 | 1.11 | 1.13 | 1.2 | 1.197 | 0.93 | 1.1255 | 1.27 | 1.04 | 1.21 | 1.05 | 1.26 |  |  |  |  | 1
| 1 | 2 | 1.12 | 1.0439 | 1.12 | 1 | 1.08 | 1.09 | 1.1 | 1.174 | 0.95 | 1.1427 | 1.26 | 1.07 | 1.17 | 1.08 | 1.25 |  |  |  |  | 31
| 1 | 3 | 1.18 | 1.0515 | 1.14 | 1.1 | 1.08 | 1.12 | 1.1 | 1.204 | 0.91 | 1.1215 | 1.27 | 1.02 | 1.14 | 1.04 | 1.24 |  |  |  |  | 16
| 1 | 4 | 1.15 | 1.0389 | 1.14 | 1 | 1.09 | 1.08 | 1.2 | 1.194 | 0.98 | 1.1688 | 1.27 | 1.04 | 1.19 | 1.02 | 1.33 |  |  |  |  | 1.1255
| 1 | 5 | 1.13 | 1.0566 | 1.11 | 1.1 | 1.1 | 1.11 | 1.2 | 1.227 | 0.96 | 1.1755 | 1.29 | 1.09 | 1.21 | 1.02 | 1.26 |  |  |  |  | 115
| 1 | 6 | 1.15 | 1.0726 | 1.14 | 1.1 | 1.09 | 1.1 | 1.2 | 1.284 | 0.97 | 1.1809 | 1.26 | 1.08 | 1.23 | 1.02 | 1.23
| Mean |  | 1.1517 | 1.0478 | 1.13 | 1.0833 | 1.0917 | 1.105 | 1.1667 | 1.2133 | 0.95 | 1.1525 | 1.27 | 1.0567 | 1.1917 | 1.0383 | 1.2617
| Median |  | 1.15 | 1.0477 | 1.135 | 1.1 | 1.09 | 1.105 | 1.2 | 1.2005 | 0.955 | 1.1557 | 1.27 | 1.055 | 1.2 | 1.03 | 1.255
| Std Dev. |  | 0.024833 | 0.016821 | 0.012649 | 0.075277 | 0.01169 | 0.018708 | 0.05164 | 0.038614 | 0.026077 | 0.026032 | 0.010954 | 0.027325 | 0.032506 | 0.024014 | 0.035449
| Rel.Std.Dev. |  | 2.156% | 1.605% | 1.119% | 6.949% | 1.071% | 1.693% | 4.426% | 3.182% | 2.745% | 2.259% | 0.8626% | 2.586% | 2.728% | 2.313% | 2.81%
| PDM3 |  | 2.325% | -6.908% | 0.4002% | -3.746% | -3.006% | -1.821% | 3.658% | 7.804% | -15.59% | 2.397% | 12.84% | -6.115% | 5.879% | -7.744% | 12.1%
| Z-Score (Absolute) |  | 0.24 | 0.9 | 0 | 0.51 | 0.42 | 0.27 | 0.4 | 0.91 | 1.97 | 0.25 | 1.53 | 0.8 | 0.67 | 1 | 1.44
| Analytical results for V in OREAS 262b (Certified Value 19.3 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 20 | 19.033 | 18.085 | 22 | 22 | 18 | 19 | 20.8556 | 19 | 21 | 18 | 18.7158 | 20 | 19 | 18 | 19 |  |  |  | 1
| 1 | 2 | 20 | 18.9652 | 18.935 | 22 | 22 | 17 | 19 | 20.2106 | 19 | 20 | 18 | 19.0275 | 20 | 19 | 18 | 19 |  |  |  | 32
| 1 | 3 | 19 | 18.8618 | 18.531 | 21 | 22 | 18 | 19 | 20.8857 | 19 | 20 | 18 | 19.1796 | 20 | 19 | 18 | 19 |  |  |  | 16
| 1 | 4 | 20 | 19.2254 | 18.858 | 23 | 23 | 17 | 19 | 20.8296 | 19 | 20 | 18 | 20.0107 | 20 | 19 | 19 | 19 |  |  |  | 19.2988
| 1 | 5 | 20 | 19.6657 | 18.492 | 21 | 22 | 18 | 19 | 20.7129 | 19 | 21 | 18 | 19.9919 | 20 | 20 | 18 | 19 |  |  |  | 116
| 1 | 6 | 19 | 19.3579 | 18.706 | 21 | 22 | 18 | 19 | 20.6695 | 20 | 20 | 18 | 19.6819 | 20 | 20 | 18 | 18
| Mean |  | 19.6667 | 19.1848 | 18.6012 | 21.6667 | 22.1667 | 17.6667 | 19 | 20.694 | 19.1667 | 20.3333 | 18 | 19.4346 | 20 | 19.3333 | 18.1667 | 18.8333
| Median |  | 20 | 19.1292 | 18.6185 | 21.5 | 22 | 18 | 19 | 20.7713 | 19 | 20 | 18 | 19.4308 | 20 | 19 | 18 | 19
| Std Dev. |  | 0.5164 | 0.29614 | 0.30714 | 0.8165 | 0.40825 | 0.5164 | 0 | 0.25135 | 0.40825 | 0.5164 | 0 | 0.53873 | 0 | 0.5164 | 0.40825 | 0.40825
| Rel.Std.Dev. |  | 2.626% | 1.544% | 1.651% | 3.768% | 1.842% | 2.923% | 0% | 1.215% | 2.13% | 2.54% | 0% | 2.772% | 0% | 2.671% | 2.247% | 2.168%
| PDM3 |  | 1.906% | -0.5902% | -3.615% | 12.27% | 14.86% | -8.457% | -1.548% | 7.23% | -0.6844% | 5.361% | -6.73% | 0.7038% | 3.634% | 0.1792% | -5.866% | -2.412%
| Z-Score (Absolute) |  | 0.39 | 0.07 | 0.63 | 2.32 | 2.8 | 1.54 | 0.25 | 1.38 | 0.09 | 1.04 | 1.21 | 0.17 | 0.71 | 0.07 | 1.05 | 0.41
| Analytical results for W in OREAS 262b (Certified Value 0.64 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.82 | < 1 | 1 | 0.5 | 0.7 | 0.45 | < 5 | 0.71 | 0.6 | 0.69 | 0.4527 | 0.8 | 0.64 | 0.81 | 0.53 |  |  |  |  | 1
| 1 | 2 | 0.8 | < 1 | < 1 | 0.5 | 0.5 | 0.46 | < 5 | 0.72 | 0.6 | 0.73 | 0.52684 | 0.8 | 0.66 | 0.79 | 0.55 |  |  |  |  | 33
| 1 | 3 | 0.83 | < 1 | 1 | 0.5 | 0.6 | 0.45 | < 5 | 0.71 | 0.6 | 0.64 | 0.46446 | 0.7 | 0.65 | 0.76 | 0.57 |  |  |  |  | 16
| 1 | 4 | 0.81 | < 1 | < 1 | 0.5 | 0.5 | 0.45 | < 5 | 0.69 | 0.6 | 0.71 | 0.54942 | 0.8 | 0.68 | 0.77 | 0.52 |  |  |  |  | 0.64274
| 1 | 5 | 0.8 | < 1 | < 1 | 0.4 | 0.5 | 0.48 | < 5 | 0.68 | 0.6 | 0.66 | 0.54614 | 0.8 | 0.64 | 0.78 | 0.52 |  |  |  |  | 117
| 1 | 6 | 0.81 | < 1 | < 1 | 0.5 | 0.5 | 0.47 | < 5 | 0.79 | 0.7 | 0.7 | 0.57013 | 0.7 | 0.65 | 0.71 | 0.54
| Mean |  | 0.81167 |  | 1 | 0.48333 | 0.55 | 0.46 |  | 0.71667 | 0.61667 | 0.68833 | 0.51828 | 0.76667 | 0.65333 | 0.77 | 0.53833
| Median |  | 0.81 |  | 1 | 0.5 | 0.5 | 0.455 |  | 0.71 | 0.6 | 0.695 | 0.53649 | 0.8 | 0.65 | 0.775 | 0.535
| Std Dev. |  | 0.01169 |  | 0 | 0.040825 | 0.083666 | 0.012649 |  | 0.038816 | 0.040825 | 0.033116 | 0.048387 | 0.05164 | 0.015055 | 0.034059 | 0.019408
| Rel.Std.Dev. |  | 1.44% |  | 0% | 8.447% | 15.21% | 2.75% |  | 5.416% | 6.62% | 4.811% | 9.336% | 6.736% | 2.304% | 4.423% | 3.605%
| PDM3 |  | 26.28% |  | 55.58% | -24.8% | -14.43% | -28.43% |  | 11.5% | -4.057% | 7.093% | -19.36% | 19.28% | 1.648% | 19.8% | -16.24%
| Z-Score (Absolute) |  | 0.69 | 0.84 | NA | NA | NA | 1.04 | 9.03 | 0.23 | NA | 0.09 | 0.75 | NA | 0.09 | 0.49 | 0.65
Results from laboratories 1.07, 1.09, 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 262b (Certified Value 5.06 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 4.79 | 5.1674 | 6 | 5.21 | 5.09 | 4.98 | 5.3067 | 4.78 | 4.05 | 5.6401 | 5.5 | 4.75 | 5.65 | 4.81 | 5.36 |  |  |  |  | 1
| 1 | 2 | 4.68 | 5.2478 | 6 | 5.59 | 4.69 | 4.87 | 5.115 | 4.77 | 4.01 | 5.6076 | 5.5 | 5.06 | 5.13 | 4.83 | 5.32 |  |  |  |  | 34
| 1 | 3 | 4.76 | 5.1273 | 6 | 5.19 | 4.92 | 4.91 | 5.32 | 5.08 | 3.92 | 5.5499 | 5.5 | 4.89 | 4.94 | 4.88 | 5.33 |  |  |  |  | 16
| 1 | 4 | 4.64 | 5.1588 | 7 | 5.46 | 4.69 | 4.92 | 5.3 | 4.75 | 4.2 | 5.8052 | 5.4 | 4.81 | 5.45 | 4.8 | 5.37 |  |  |  |  | 5.0566
| 1 | 5 | 4.72 | 5.2755 | 6 | 5.33 | 5.02 | 5.02 | 5.2533 | 4.82 | 4.15 | 5.7615 | 5.6 | 5.08 | 5.08 | 4.61 | 5.32 |  |  |  |  | 118
| 1 | 6 | 4.71 | 5.1453 | 6 | 5.28 | 5.07 | 4.85 | 5.32 | 4.77 | 4.21 | 5.7523 | 5.5 | 5.03 | 5.68 | 4.64 | 5.3
| Mean |  | 4.7167 | 5.187 | 6.1667 | 5.3433 | 4.9133 | 4.925 | 5.2692 | 4.8283 | 4.09 | 5.6861 | 5.5 | 4.9367 | 5.3217 | 4.7617 | 5.3333
| Median |  | 4.715 | 5.1631 | 6 | 5.305 | 4.97 | 4.915 | 5.3033 | 4.775 | 4.1 | 5.6962 | 5.5 | 4.96 | 5.29 | 4.805 | 5.325
| Std Dev. |  | 0.053914 | 0.059995 | 0.40825 | 0.15488 | 0.18272 | 0.06473 | 0.079427 | 0.12545 | 0.11576 | 0.10107 | 0.063246 | 0.13967 | 0.31416 | 0.1098 | 0.026583
| Rel.Std.Dev. |  | 1.143% | 1.157% | 6.62% | 2.898% | 3.719% | 1.314% | 1.507% | 2.598% | 2.83% | 1.778% | 1.15% | 2.829% | 5.903% | 2.306% | 0.4984%
| PDM3 |  | -6.723% | 2.579% | 21.95% | 5.67% | -2.834% | -2.603% | 4.203% | -4.515% | -19.12% | 12.45% | 8.768% | -2.372% | 5.241% | -5.833% | 5.472%
| Z-Score (Absolute) |  | 0.88 | 0.32 | NA | 0.71 | 0.38 | 0.35 | 0.53 | 0.59 | 2.47 | 1.58 | 1.11 | 0.32 | 0.66 | 0.76 | 0.69
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 262b (Certified Value 0.44 ppm) | [20 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09; 1.09: 1.17; 1.10: 1.22 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g; 1.09: 0.2g; 1.10: 0.25g |  |  |  |  | 2
| 1 | 1 | 1.01: 0.417; 1.03: 0.41501; 1.07: 0.5; 1.09: 0.4; 1.10: 0.5 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.405; 1.03: 0.39388; 1.07: 0.4; 1.09: 0.4; 1.10: 0.5 |  |  |  |  | 35
| 1 | 3 | 1.01: 0.417; 1.03: 0.3916; 1.07: 0.5; 1.09: 0.4; 1.10: 0.5 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.403; 1.03: 0.39585; 1.07: 0.4; 1.09: 0.5; 1.10: 0.5 |  |  |  |  | 0.43635
| 1 | 5 | 1.01: 0.4; 1.03: 0.4245; 1.07: 0.4; 1.09: 0.5; 1.10: 0.5 |  |  |  |  | 119
| 1 | 6 | 1.01: 0.405; 1.03: 0.42261; 1.07: 0.4; 1.09: 0.4; 1.10: 0.5
| Mean |  | 1.01: 0.40783; 1.03: 0.40724; 1.07: 0.43333; 1.09: 0.43333; 1.10: 0.5
| Median |  | 1.01: 0.405; 1.03: 0.40543; 1.07: 0.4; 1.09: 0.4; 1.10: 0.5
| Std Dev. |  | 1.01: 0.0073326; 1.03: 0.015146; 1.07: 0.05164; 1.09: 0.05164; 1.10: 0
| Rel.Std.Dev. |  | 1.01: 1.798%; 1.03: 3.719%; 1.07: 11.92%; 1.09: 11.92%; 1.10: 0%
| PDM3 |  | 1.01: -6.535%; 1.03: -6.671%; 1.07: -0.6909%; 1.09: -0.6909%; 1.10: 14.59%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.69; 1.07: 0; 1.09: 0; 1.10: 1.76
| Analytical results for Zn in OREAS 262b (Certified Value 28.5 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 28 | 39.0528 | 29.1627 | 21.984 | 34 | 34 | 39.7 | 28 | 26.4333 | 28 | 32 | 22.521 | 26 | 32.4801 | 34.6 | 29 | 27 | 27 | 27 | 1
| 1 | 2 | 27 | 40.9752 | 29.5577 | 24.137 | 32 | 34 | 38.8 | 28 | 26.52 | 28 | 29 | 23.325 | 26 | 32.5834 | 35.3 | 29 | 26 | 26 | 28 | 36
| 1 | 3 | 27 | 38.9232 | 29.075 | 23.177 | 29 | 34 | 39.2 | 28 | 26.9133 | 29 | 29 | 22.989 | 26 | 34.0611 | 34.6 | 29 | 27 | 27 | 26 | 16
| 1 | 4 | 27 | 38.6076 | 29.0017 | 23.519 | 34 | 34 | 39.7 | 28 | 26.73 | 29 | 29 | 22.996 | 26 | 34.0136 | 35.4 | 29 | 27 | 26 | 27 | 28.4996
| 1 | 5 | 27 | 38.4652 | 29.9065 | 22.825 | 33 | 33 | 39.6 | 28 | 28.57 | 29 | 31 | 22.996 | 26 | 33.2075 | 35.2 | 30 | 26 | 27 | 25 | 120
| 1 | 6 | 28 | 38.3376 | 28.9496 | 23.921 | 32 | 33 | 39.6 | 29 | 26.2933 | 27 | 29 | 22.826 | 26 | 33.2778 | 34.5 | 31 | 26 | 26 | 26
| Mean |  | 27.3333 | 39.0603 | 29.2755 | 23.2605 | 32.3333 | 33.6667 | 39.4333 | 28.1667 | 26.91 | 28.3333 | 29.8333 | 22.9422 | 26 | 33.2706 | 34.9333 | 29.5 | 26.5 | 26.5 | 26.5
| Median |  | 27 | 38.7654 | 29.1189 | 23.348 | 32.5 | 34 | 39.6 | 28 | 26.625 | 28.5 | 29 | 22.9925 | 26 | 33.2427 | 34.9 | 29 | 26.5 | 26.5 | 26.5
| Std Dev. |  | 0.5164 | 0.97643 | 0.37735 | 0.78707 | 1.8619 | 0.5164 | 0.36148 | 0.40825 | 0.84233 | 0.8165 | 1.3292 | 0.26284 | 0 | 0.675 | 0.40825 | 0.83666 | 0.54772 | 0.54772 | 1.0488
| Rel.Std.Dev. |  | 1.889% | 2.5% | 1.289% | 3.384% | 5.758% | 1.534% | 0.9167% | 1.449% | 3.13% | 2.882% | 4.455% | 1.146% | 0% | 2.029% | 1.169% | 2.836% | 2.067% | 2.067% | 3.958%
| PDM3 |  | -4.092% | 37.06% | 2.723% | -18.38% | 13.45% | 18.13% | 38.36% | -1.168% | -5.578% | -0.5835% | 4.68% | -19.5% | -8.771% | 16.74% | 22.57% | 3.51% | -7.016% | -7.016% | -7.016%
| Z-Score (Absolute) |  | 0.37 | 3.95 | 0.35 | 1.87 | 1.47 | 1.96 | 4.08 | 0.06 | 0.52 | 0 | 0.55 | 1.98 | 0.86 | 1.82 | 2.43 | 0.43 | 0.67 | 0.67 | 0.67
| Analytical results for Zr in OREAS 262b (Certified Value 13.4 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 14.1 | 14.3871 | 51 | 10 | 10.5 | 8 | 13.2925 | 14.7 | 14.1 | 13.5483 | 11.4 | 12.8 | 16.9 | 13.5 | 16 |  |  |  |  | 1
| 1 | 2 | 13.6 | 14.7143 | 46 | 9.9 | 9.2 | 9 | 13.3127 | 15.2 | 14.1 | 14.055 | 10.6 | 13.6 | 15.7 | 13.5 | 16.4 |  |  |  |  | 37
| 1 | 3 | 13.8 | 14.5338 | 45 | 9.9 | 10 | 8 | 13.2902 | 15.2 | 13.6 | 13.3311 | 11.6 | 13 | 15.5 | 13.8 | 15.9 |  |  |  |  | 16
| 1 | 4 | 13.7 | 14.6039 | 69 | 9.7 | 8.7 | 9 | 13.2124 | 15 | 14.3 | 13.94 | 11.6 | 13.6 | 16.5 | 13.8 | 15.8 |  |  |  |  | 13.3798
| 1 | 5 | 13.7 | 14.7431 | 61 | 9.8 | 9.1 | 9 | 13.0283 | 14.8 | 13.8 | 13.9571 | 12.1 | 13.4 | 15.9 | 12.8 | 14.6 |  |  |  |  | 121
| 1 | 6 | 13.6 | 14.9887 | 49 | 9.7 | 9.6 | 10 | 13.0797 | 15 | 14.4 | 13.8501 | 9.6 | 13.5 | 17.1 | 13.2 | 15.4
| Mean |  | 13.75 | 14.6618 | 53.5 | 9.8333 | 9.5167 | 8.8333 | 13.2026 | 14.9833 | 14.05 | 13.7803 | 11.15 | 13.3167 | 16.2667 | 13.4333 | 15.6833
| Median |  | 13.7 | 14.6591 | 50 | 9.85 | 9.4 | 9 | 13.2513 | 15 | 14.1 | 13.8951 | 11.5 | 13.45 | 16.2 | 13.5 | 15.85
| Std Dev. |  | 0.18708 | 0.20565 | 9.5026 | 0.12111 | 0.65549 | 0.75277 | 0.12122 | 0.20412 | 0.30166 | 0.28028 | 0.90277 | 0.33714 | 0.66232 | 0.38297 | 0.62102
| Rel.Std.Dev. |  | 1.361% | 1.403% | 17.76% | 1.232% | 6.888% | 8.522% | 0.9181% | 1.362% | 2.147% | 2.034% | 8.097% | 2.532% | 4.072% | 2.851% | 3.96%
| PDM3 |  | 2.766% | 9.581% | 299.9% | -26.51% | -28.87% | -33.98% | -1.325% | 11.98% | 5.009% | 2.993% | -16.67% | -0.4722% | 21.58% | 0.3997% | 17.22%
| Z-Score (Absolute) |  | 0.09 | 0.59 | NA | 2.06 | 2.23 | 2.61 | 0.21 | 0.76 | 0.25 | 0.1 | 1.34 | 0.15 | 1.47 | 0.09 | 1.15
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C533: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C551: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C569: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C606: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C624: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C642: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C660: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C987: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1042: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 262b (Indicative Value 9.58 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 9.58 | 1
| 1 | 2 | 9.58 | 12
| Mean |  | 9.58 | 16
| Median |  | 9.58 | 9.58
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 262b (Indicative Value 0.49 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.49 | 1
| 1 | 2 | 0.49 | 13
| Mean |  | 0.49 | 16
| Median |  | 0.49 | 0.49
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 262b (Indicative Value 3.51 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.51 | 1
| 1 | 2 | 3.51 | 14
| Mean |  | 3.51 | 16
| Median |  | 3.51 | 3.51
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 262b (Indicative Value 1.97 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.97 | 1
| 1 | 2 | 1.97 | 15
| Mean |  | 1.97 | 16
| Median |  | 1.97 | 1.97
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 262b (Indicative Value 0.945 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.94 | 12
| Mean |  | 0.945 | 16
| Median |  | 0.945 | 0.945
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 262b (Indicative Value 0.023 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.023 | 1
| 1 | 2 | 0.023 | 13
| Mean |  | 0.023 | 16
| Median |  | 0.023 | 0.023
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 262b (Indicative Value 0.26 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.26 | 1
| 1 | 2 | 0.26 | 14
| Mean |  | 0.26 | 16
| Median |  | 0.26 | 0.26
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 262b (Indicative Value 0.057 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.056 | 1
| 1 | 2 | 0.058 | 15
| Mean |  | 0.057 | 16
| Median |  | 0.057 | 0.057
| Std Dev. |  | 0.0014142 | 21
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 262b (Indicative Value 80.03 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 80.06 | 1
| 1 | 2 | 79.99 | 12
| Mean |  | 80.025 | 16
| Median |  | 80.025 | 80.025
| Std Dev. |  | 0.049497 | 18
| Rel.Std.Dev. |  | 0.06185%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 262b (Indicative Value 0.092 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.093 | 1
| 1 | 2 | 0.09 | 13
| Mean |  | 0.0915 | 16
| Median |  | 0.0915 | 0.0915
| Std Dev. |  | 0.0021213 | 19
| Rel.Std.Dev. |  | 2.318%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 262b (Indicative Value 0.49 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.491 | 1
| 1 | 2 | 0.488 | 14
| Mean |  | 0.4895 | 16
| Median |  | 0.4895 | 0.4895
| Std Dev. |  | 0.0021213 | 20
| Rel.Std.Dev. |  | 0.4334%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 262b (Indicative Value 2.45 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.44 | 1
| 1 | 2 | 2.45 | 17
| Mean |  | 2.445 | 16
| Median |  | 2.445 | 2.445
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 0.2892%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 262b (Indicative Value 0.145 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.14 | 1
| 1 | 2 | 0.15 | 19
| Mean |  | 0.145 | 16
| Median |  | 0.145 | 0.145
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 262b (Indicative Value 0.02 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.02 | 19
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.02
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 262b (Indicative Value 0.15 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.2 | 21
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0.070711 | 27
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 262b (Indicative Value 29.7 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 29.6 | 1
| 1 | 2 | 29.8 | 22
| Mean |  | 29.7 | 16
| Median |  | 29.7 | 29.7
| Std Dev. |  | 0.14142 | 28
| Rel.Std.Dev. |  | 0.4762%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 262b (Indicative Value 325 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 321 | 1
| 1 | 2 | 329 | 23
| Mean |  | 325 | 16
| Median |  | 325 | 325
| Std Dev. |  | 5.6569 | 29
| Rel.Std.Dev. |  | 1.741%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 262b (Indicative Value 1.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 1.8 | 24
| Mean |  | 1.9 | 16
| Median |  | 1.9 | 1.9
| Std Dev. |  | 0.14142 | 30
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 262b (Indicative Value 0.79 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.82 | 1
| 1 | 2 | 0.76 | 25
| Mean |  | 0.79 | 16
| Median |  | 0.79 | 0.79
| Std Dev. |  | 0.042426 | 31
| Rel.Std.Dev. |  | 5.37%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 262b (Indicative Value 0.13 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | 0.2 | 26
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.125
| Std Dev. |  |  | 32
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 262b (Indicative Value 66 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 65.8 | 1
| 1 | 2 | 66.1 | 27
| Mean |  | 65.95 | 16
| Median |  | 65.95 | 65.95
| Std Dev. |  | 0.21213 | 33
| Rel.Std.Dev. |  | 0.3217%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 262b (Indicative Value 6.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 6.4 | 28
| Mean |  | 6.1 | 16
| Median |  | 6.1 | 6.1
| Std Dev. |  | 0.42426 | 34
| Rel.Std.Dev. |  | 6.955%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 262b (Indicative Value 57 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 55 | 1
| 1 | 2 | 59 | 29
| Mean |  | 57 | 16
| Median |  | 57 | 57
| Std Dev. |  | 2.8284 | 35
| Rel.Std.Dev. |  | 4.962%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 262b (Indicative Value 4.85 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.82 | 1
| 1 | 2 | 4.88 | 30
| Mean |  | 4.85 | 16
| Median |  | 4.85 | 4.85
| Std Dev. |  | 0.042426 | 36
| Rel.Std.Dev. |  | 0.8748%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 262b (Indicative Value 71 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72 | 1
| 1 | 2 | 70 | 31
| Mean |  | 71 | 16
| Median |  | 71 | 71
| Std Dev. |  | 1.4142 | 37
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 262b (Indicative Value 3.83 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.77 | 1
| 1 | 2 | 3.88 | 5
| Mean |  | 3.825 | 16
| Median |  | 3.825 | 3.825
| Std Dev. |  | 0.077782 | 38
| Rel.Std.Dev. |  | 2.034%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 262b (Indicative Value 2.24 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.27 | 6
| Mean |  | 2.235 | 16
| Median |  | 2.235 | 2.235
| Std Dev. |  | 0.049497 | 39
| Rel.Std.Dev. |  | 2.215%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 262b (Indicative Value 0.79 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.78 | 1
| 1 | 2 | 0.79 | 7
| Mean |  | 0.785 | 16
| Median |  | 0.785 | 0.785
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.9008%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 262b (Indicative Value 11.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.8 | 1
| 1 | 2 | 11.8 | 35
| Mean |  | 11.8 | 16
| Median |  | 11.8 | 11.8
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 262b (Indicative Value 3.89 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.79 | 1
| 1 | 2 | 3.98 | 8
| Mean |  | 3.885 | 16
| Median |  | 3.885 | 3.885
| Std Dev. |  | 0.13435 | 42
| Rel.Std.Dev. |  | 3.458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 262b (Indicative Value 1.25 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.25 | 3
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 262b (Indicative Value 6.74 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.83 | 1
| 1 | 2 | 6.65 | 21
| Mean |  | 6.74 | 16
| Median |  | 6.74 | 6.74
| Std Dev. |  | 0.12728 | 27
| Rel.Std.Dev. |  | 1.888%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 262b (Indicative Value 0.78 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.78 | 1
| 1 | 2 | 0.77 | 10
| Mean |  | 0.775 | 16
| Median |  | 0.775 | 0.775
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.9124%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 262b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 23
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 262b (Indicative Value 34.2 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 34.5 | 1
| 1 | 2 | 33.9 | 24
| Mean |  | 34.2 | 16
| Median |  | 34.2 | 34.2
| Std Dev. |  | 0.42426 | 30
| Rel.Std.Dev. |  | 1.241%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 262b (Indicative Value 0.34 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.33 | 5
| Mean |  | 0.335 | 16
| Median |  | 0.335 | 0.335
| Std Dev. |  | 0.0070711 | 31
| Rel.Std.Dev. |  | 2.111%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 262b (Indicative Value 0.018 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0184 | 1
| 1 | 2 | 0.0184 | 26
| Mean |  | 0.0184 | 16
| Median |  | 0.0184 | 0.0184
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 262b (Indicative Value 1.7 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.6 | 27
| Mean |  | 1.7 | 16
| Median |  | 1.7 | 1.7
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 262b (Indicative Value 10.4 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.3 | 1
| 1 | 2 | 10.5 | 7
| Mean |  | 10.4 | 16
| Median |  | 10.4 | 10.4
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 1.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 262b (Indicative Value 26.6 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.6 | 1
| 1 | 2 | 26.6 | 8
| Mean |  | 26.6 | 16
| Median |  | 26.6 | 26.6
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 262b (Indicative Value 22 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 22 | 1
| 1 | 2 | 22 | 30
| Mean |  | 22 | 16
| Median |  | 22 | 22
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 262b (Indicative Value 20 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 31
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 262b (Indicative Value 7.66 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.69 | 1
| 1 | 2 | 7.62 | 10
| Mean |  | 7.655 | 16
| Median |  | 7.655 | 7.655
| Std Dev. |  | 0.049497 | 38
| Rel.Std.Dev. |  | 0.6466%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 262b (Indicative Value 104 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 104 | 1
| 1 | 2 | 104 | 33
| Mean |  | 104 | 16
| Median |  | 104 | 104
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 262b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 34
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 40
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 262b (Indicative Value 6.25 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.2 | 1
| 1 | 2 | 6.3 | 35
| Mean |  | 6.25 | 16
| Median |  | 6.25 | 6.25
| Std Dev. |  | 0.070711 | 41
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 262b (Indicative Value 6.8 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.1 | 1
| 1 | 2 | 6.5 | 36
| Mean |  | 6.8 | 16
| Median |  | 6.8 | 6.8
| Std Dev. |  | 0.42426 | 42
| Rel.Std.Dev. |  | 6.239%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 262b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 37
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 43
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 262b (Indicative Value 4.68 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.69 | 1
| 1 | 2 | 4.67 | 7
| Mean |  | 4.68 | 16
| Median |  | 4.68 | 4.68
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 0.3022%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 262b (Indicative Value 9.9 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.8 | 1
| 1 | 2 | 10 | 22
| Mean |  | 9.9 | 16
| Median |  | 9.9 | 9.9
| Std Dev. |  | 0.14142 | 28
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 262b (Indicative Value 41.9 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 41.4 | 1
| 1 | 2 | 42.4 | 23
| Mean |  | 41.9 | 16
| Median |  | 41.9 | 41.9
| Std Dev. |  | 0.70711 | 29
| Rel.Std.Dev. |  | 1.688%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 262b (Indicative Value 0.93 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.93 | 1
| 1 | 2 | 0.93 | 24
| Mean |  | 0.93 | 16
| Median |  | 0.93 | 0.93
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 262b (Indicative Value 0.63 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.62 | 1
| 1 | 2 | 0.63 | 25
| Mean |  | 0.625 | 16
| Median |  | 0.625 | 0.625
| Std Dev. |  | 0.0070711 | 31
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 262b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 26
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 32
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 262b (Indicative Value 12.9 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 13.2 | 27
| Mean |  | 12.9 | 16
| Median |  | 12.9 | 12.9
| Std Dev. |  | 0.42426 | 33
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 262b (Indicative Value 0.29 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.29 | 28
| Mean |  | 0.29 | 16
| Median |  | 0.29 | 0.29
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 262b (Indicative Value 0.8 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.8 | 29
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 262b (Indicative Value 0.33 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.33 | 8
| Mean |  | 0.325 | 16
| Median |  | 0.325 | 0.325
| Std Dev. |  | 0.0070711 | 36
| Rel.Std.Dev. |  | 2.176%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 262b (Indicative Value 2.19 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.17 | 31
| Mean |  | 2.185 | 16
| Median |  | 2.185 | 2.185
| Std Dev. |  | 0.021213 | 37
| Rel.Std.Dev. |  | 0.9709%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 262b (Indicative Value 54 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 53.5 | 1
| 1 | 2 | 54.5 | 32
| Mean |  | 54 | 16
| Median |  | 54 | 54
| Std Dev. |  | 0.70711 | 38
| Rel.Std.Dev. |  | 1.309%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 262b (Indicative Value 3.25 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 3.5 | 33
| Mean |  | 3.25 | 16
| Median |  | 3.25 | 3.25
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 262b (Indicative Value 20.7 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.7 | 1
| 1 | 2 | 20.6 | 34
| Mean |  | 20.65 | 16
| Median |  | 20.65 | 20.65
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 0.3424%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 262b (Indicative Value 2.23 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.27 | 1
| 1 | 2 | 2.18 | 35
| Mean |  | 2.225 | 16
| Median |  | 2.225 | 2.225
| Std Dev. |  | 0.06364 | 41
| Rel.Std.Dev. |  | 2.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 262b (Indicative Value 47.5 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 45 | 36
| Mean |  | 47.5 | 16
| Median |  | 47.5 | 47.5
| Std Dev. |  | 3.5355 | 42
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 262b (Indicative Value 234 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 235 | 1
| 1 | 2 | 232 | 37
| Mean |  | 233.5 | 16
| Median |  | 233.5 | 233.5
| Std Dev. |  | 2.1213 | 43
| Rel.Std.Dev. |  | 0.9085%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.