## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 466
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 0.32234 | 0.25848 | 0.3862 | 0.28736 | 0.35732
Al, Aluminium (wt.%) | 7.8995 | 7.68 | 8.119 | 7.7588 | 8.0401
As, Arsenic (ppm) | 481.4941 | 463.1677 | 499.8204 | 471.9419 | 491.0462
Ba, Barium (ppm) | 701.7594 | 678.678 | 724.8408 | 689.0536 | 714.4652
Be, Beryllium (ppm) | 2.9045 | 2.7713 | 3.0376 | 2.7956 | 3.0133
Bi, Bismuth (ppm) | 22.0859 | 20.5026 | 23.6692 | 20.9527 | 23.2192
Ca, Calcium (wt.%) | 0.22587 | 0.21349 | 0.23826 | 0.21942 | 0.23233
Ce, Cerium (ppm) | 339.3862 | 324.6545 | 354.1178 | 332.039 | 346.7333
Co, Cobalt (ppm) | 2.9229 | 2.7447 | 3.1011 | 2.8057 | 3.0401
Cr, Chromium (ppm) | 131.7938 | 122.9222 | 140.6654 | 127.0086 | 136.579
Cs, Caesium (ppm) | 10.7543 | 10.2609 | 11.2477 | 10.3905 | 11.1182
Cu, Copper (ppm) | 34.1434 | 32.7163 | 35.5706 | 33.137 | 35.1499
Dy, Dysprosium (ppm) | 5.5683 | 5.2014 | 5.9353 | 5.3737 | 5.7629
Er, Erbium (ppm) | 2.1932 | 1.9658 | 2.4205 | 2.1039 | 2.2824
Eu, Europium (ppm) | 4.2557 | 3.9625 | 4.5489 | 4.1319 | 4.3795
Fe, Iron (wt.%) | 4.9673 | 4.8079 | 5.1267 | 4.855 | 5.0796
Ga, Gallium (ppm) | 22.3564 | 21.3477 | 23.365 | 21.8205 | 22.8922
Gd, Gadolinium (ppm) | 11.4013 | 10.5156 | 12.287 | 11.0101 | 11.7925
Ge, Germanium (ppm) | 0.42109 | 0.31795 | 0.52423 | 0.38517 | 0.45701
Hf, Hafnium (ppm) | 4.6995 | 4.3914 | 5.0077 | 4.5153 | 4.8838
Ho, Holmium (ppm) | 0.86995 | 0.78489 | 0.95502 | 0.83258 | 0.90732
In, Indium (ppm) | 0.12086 | 0.10588 | 0.13584 | 0.11276 | 0.12896
K, Potassium (wt.%) | 2.3934 | 2.334 | 2.4528 | 2.338 | 2.4488
La, Lanthanum (ppm) | 229.0675 | 215.0231 | 243.1119 | 221.1509 | 236.9841
Li, Lithium (ppm) | 43.8747 | 42.4244 | 45.3251 | 42.6807 | 45.0688
Lu, Lutetium (ppm) | 0.28168 | 0.24767 | 0.3157 | 0.26597 | 0.29739
Mg, Magnesium (wt.%) | 0.41156 | 0.38721 | 0.43591 | 0.39922 | 0.4239
Mn, Manganese (wt.%) | 0.0091496 | 0.008714 | 0.0095852 | 0.0088453 | 0.0094539
Mo, Molybdenum (ppm) | 6.4168 | 6.0748 | 6.7587 | 6.1483 | 6.6852
Na, Sodium (wt.%) | 0.1638 | 0.15269 | 0.1749 | 0.1583 | 0.1693
Nb, Niobium (ppm) | 109.4805 | 103.6719 | 115.289 | 105.3922 | 113.5687
Nd, Neodymium (ppm) | 152.0785 | 145.5874 | 158.5696 | 147.9856 | 156.1714
Ni, Nickel (ppm) | 19.0014 | 17.9183 | 20.0846 | 18.1731 | 19.8297
P, Phosphorus (wt.%) | 0.054981 | 0.05247 | 0.057493 | 0.053543 | 0.05642
Pb, Lead (ppm) | 22.988 | 21.3526 | 24.6234 | 21.8603 | 24.1158
Pr, Praseodymium (ppm) | 44.1153 | 41.9743 | 46.2562 | 42.8649 | 45.3657
Rb, Rubidium (ppm) | 140.2327 | 133.7222 | 146.7432 | 135.4431 | 145.0223
S, Sulphur (wt.%) | 0.011701 | 0.0091004 | 0.014302 | IND | IND
Sb, Antimony (ppm) | 5.4451 | 5.2395 | 5.6506 | 5.1972 | 5.693
Sc, Scandium (ppm) | 18.8402 | 17.8529 | 19.8275 | 18.2156 | 19.4649
Sm, Samarium (ppm) | 20.8147 | 19.4788 | 22.1505 | 19.7571 | 21.8722
Sn, Tin (ppm) | 17.6454 | 16.4983 | 18.7925 | 16.8972 | 18.3936
Sr, Strontium (ppm) | 107.0192 | 102.6605 | 111.3778 | 104.2112 | 109.8271
Ta, Tantalum (ppm) | 2.5637 | 2.3698 | 2.7575 | 2.4287 | 2.6986
Tb, Terbium (ppm) | 1.2646 | 1.1559 | 1.3732 | 1.2365 | 1.2927
Te, Tellurium (ppm) | 0.30302 | 0.22728 | 0.37875 | 0.25247 | 0.35356
Th, Thorium (ppm) | 32.4156 | 30.5507 | 34.2806 | 31.2094 | 33.6218
Ti, Titanium (wt.%) | 0.44801 | 0.412 | 0.48402 | 0.42545 | 0.47058
Tl, Thallium (ppm) | 0.78802 | 0.72409 | 0.85195 | 0.75684 | 0.81919
Tm, Thulium (ppm) | 0.27431 | 0.24479 | 0.30383 | 0.25709 | 0.29152
U, Uranium (ppm) | 3.0374 | 2.8558 | 3.219 | 2.9212 | 3.1536
V, Vanadium (ppm) | 127.5664 | 122.7384 | 132.3945 | 124.154 | 130.9788
W, Tungsten (ppm) | 90.1577 | 83.395 | 96.9205 | 85.9861 | 94.3293
Y, Yttrium (ppm) | 20.2983 | 18.998 | 21.5985 | 19.6393 | 20.9573
Yb, Ytterbium (ppm) | 1.8252 | 1.6748 | 1.9756 | 1.734 | 1.9164
Zn, Zinc (ppm) | 41.8345 | 39.8741 | 43.7948 | 40.054 | 43.615
Zr, Zirconium (ppm) | 161.3336 | 153.4272 | 169.2399 | 156.4199 | 166.2472
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 8.345 | 8.1818 | 8.5082 | 8.2443 | 8.4456
Al2O3, Aluminium(III) oxide (wt.%) | 15.7678 | 15.4594 | 16.0762 | 15.5776 | 15.958
As, Arsenic (ppm) | 472.25 | 430.9487 | 513.5513 | 453.1556 | 491.3444
B, Boron (wt.%) | 0.1383 | 0.12079 | 0.15581 | 0.13206 | 0.14454
Ba, Barium (ppm) | 718.8531 | 692.1861 | 745.5202 | 706.0069 | 731.6994
BaO, Barium oxide (ppm) | 802.5995 | 772.8258 | 832.3733 | 788.2567 | 816.9424
Be, Beryllium (ppm) | 3.1236 | 2.2768 | 3.9704 | 2.5497 | 3.6976
Bi, Bismuth (ppm) | 21.9567 | 19.3256 | 24.5877 | 19.4779 | 24.4354
Ca, Calcium (wt.%) | 0.24507 | 0.23123 | 0.25892 | 0.23333 | 0.25682
CaO, Calcium oxide (wt.%) | 0.34291 | 0.32354 | 0.36228 | 0.32647 | 0.35934
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (wt.%) | 0.035574 | 0.033629 | 0.037519 | 0.034662 | 0.036486
CeO2, Cerium(IV) oxide (wt.%) | 0.043699 | 0.04131 | 0.046088 | 0.042579 | 0.044819
Co, Cobalt (ppm) | 3.2566 | 2.5943 | 3.9189 | 2.8018 | 3.7114
Cr, Chromium (ppm) | 170.2198 | 148.2574 | 192.1821 | 153.4662 | 186.9733
Cr2O3, Chromium(III) oxide (ppm) | 248.7762 | 216.6782 | 280.8742 | 224.2909 | 273.2615
Cs, Caesium (ppm) | 10.8718 | 10.2026 | 11.5411 | 10.3844 | 11.3593
Cu, Copper (ppm) | 35.5507 | 30.4864 | 40.6149 | 28.3735 | 42.7278
Dy, Dysprosium (ppm) | 7.6412 | 6.9586 | 8.3238 | 7.2546 | 8.0279
Dy2O3, Dysprosium(III) oxide (ppm) | 8.7208 | 7.9643 | 9.4773 | 8.2795 | 9.162
Er, Erbium (ppm) | 3.6537 | 3.287 | 4.0204 | 3.4577 | 3.8497
Er2O3, Erbium(III) oxide (ppm) | 4.2414 | 3.789 | 4.6937 | 4.0139 | 4.4689
Eu, Europium (ppm) | 4.4164 | 4.0898 | 4.7429 | 4.1936 | 4.6392
Eu2O3, Europium(III) oxide (ppm) | 5.1137 | 4.7356 | 5.4918 | 4.8558 | 5.3717
Fe, Iron (wt.%) | 5.1729 | 4.977 | 5.3689 | 5.0936 | 5.2523
Fe2O3, Iron(III) oxide (wt.%) | 7.3957 | 7.1156 | 7.6759 | 7.2823 | 7.5092
Ga, Gallium (ppm) | 22.2292 | 20.2465 | 24.2119 | 20.8519 | 23.6066
Gd, Gadolinium (ppm) | 12.1392 | 11.3566 | 12.9219 | 11.683 | 12.5955
Gd2O3, Gadolinium(III) oxide (ppm) | 13.9743 | 13.151 | 14.7976 | 13.4491 | 14.4995
Hf, Hafnium (ppm) | 7.5001 | 6.7327 | 8.2675 | 7.044 | 7.9562
HfO2, Hafnium dioxide (ppm) | 8.8449 | 7.9399 | 9.7499 | 8.3069 | 9.3828
Ho, Holmium (ppm) | 1.3651 | 1.2493 | 1.4809 | 1.2922 | 1.4381
Ho2O3, Holmium(III) oxide (ppm) | 1.5638 | 1.4311 | 1.6964 | 1.4803 | 1.6473
K, Potassium (wt.%) | 2.4058 | 2.2915 | 2.5201 | 2.3647 | 2.4469
K2O, Potassium oxide (wt.%) | 2.898 | 2.7604 | 3.0357 | 2.8485 | 2.9476
La, Lanthanum (ppm) | 237.9626 | 224.9919 | 250.9334 | 231.2768 | 244.6484
La2O3, Lanthanum(III) oxide (ppm) | 279.0826 | 263.8705 | 294.2946 | 271.2415 | 286.9237
Li, Lithium (ppm) | 44.396 | 41.4317 | 47.3604 | 42.2743 | 46.5177
Lu, Lutetium (ppm) | 0.46916 | 0.40805 | 0.53027 | 0.41761 | 0.52072
Lu2O3, Lutetium(III) oxide (ppm) | 0.53004 | 0.44441 | 0.61567 | 0.47179 | 0.58828
Mg, Magnesium (wt.%) | 0.48556 | 0.45706 | 0.51406 | 0.47424 | 0.49687
MgO, Magnesium oxide (wt.%) | 0.80515 | 0.75789 | 0.85241 | 0.78639 | 0.82391
Mn, Manganese (wt.%) | 0.010021 | 0.0094096 | 0.010632 | 0.0096105 | 0.010431
MnO, Manganese oxide (wt.%) | 0.012939 | 0.01215 | 0.013728 | 0.012409 | 0.013469
Mo, Molybdenum (ppm) | 6.7233 | 5.5743 | 7.8724 | IND | IND
Nb, Niobium (ppm) | 112.9384 | 106.9516 | 118.9251 | 109.6665 | 116.2102
Nd, Neodymium (ppm) | 155.156 | 147.4288 | 162.8832 | 150.923 | 159.3889
Nd2O3, Neodymium(III) oxide (ppm) | 180.9739 | 171.9609 | 189.987 | 176.0366 | 185.9113
Ni, Nickel (ppm) | 24.6333 | 11.6086 | 37.6581 | IND | IND
P, Phosphorus (wt.%) | 0.056036 | 0.048387 | 0.063685 | 0.050295 | 0.061777
P2O5, Phosphorus(V) oxide (wt.%) | 0.12841 | 0.11088 | 0.14594 | 0.11526 | 0.14157
Pr, Praseodymium (ppm) | 44.3622 | 41.8072 | 46.9172 | 43.0146 | 45.7098
Pr6O11, Praseodymium(III,IV) oxide (ppm) | 53.5984 | 50.5115 | 56.6854 | 51.9702 | 55.2266
Rb, Rubidium (ppm) | 137.7093 | 129.6601 | 145.7586 | 132.3253 | 143.0933
Sc, Scandium (ppm) | 19.1837 | 15.4407 | 22.9267 | 17.7716 | 20.5958
Si, Silicon (wt.%) | 30.6954 | 28.7226 | 32.6682 | 30.1009 | 31.29
SiO2, Silicon dioxide (wt.%) | 65.6636 | 61.4434 | 69.8839 | 64.3918 | 66.9355
Sm, Samarium (ppm) | 20.9838 | 19.5648 | 22.4029 | 20.027 | 21.9407
Sm2O3, Samarium(III) oxide (ppm) | 24.3329 | 22.6873 | 25.9784 | 23.2233 | 25.4424
Sn, Tin (ppm) | 20.4984 | 19.2319 | 21.7649 | 18.8456 | 22.1511
Sr, Strontium (ppm) | 109.8755 | 102.8282 | 116.9229 | 105.8036 | 113.9475
Ta, Tantalum (ppm) | 2.9421 | 2.6933 | 3.1908 | 2.7108 | 3.1733
Tb, Terbium (ppm) | 1.4826 | 1.3567 | 1.6085 | 1.4136 | 1.5515
Tb4O7, Terbium(III,IV) oxide (ppm) | 1.7438 | 1.5957 | 1.8919 | 1.6627 | 1.8249
Th, Thorium (ppm) | 31.9164 | 30.295 | 33.5379 | 31.091 | 32.7418
Ti, Titanium (wt.%) | 0.58653 | 0.56623 | 0.60683 | 0.57243 | 0.60062
TiO2, Titanium dioxide (wt.%) | 0.97839 | 0.94452 | 1.0122 | 0.95488 | 1.0019
Tm, Thulium (ppm) | 0.50597 | 0.44177 | 0.57017 | 0.46847 | 0.54346
Tm2O3, Thulium(III) oxide (ppm) | 0.57612 | 0.5049 | 0.64734 | 0.53343 | 0.61881
U, Uranium (ppm) | 3.4449 | 3.1494 | 3.7404 | 3.3103 | 3.5795
V, Vanadium (ppm) | 140.0062 | 131.631 | 148.3814 | 135.932 | 144.0804
V2O5, Vanadium(V) oxide (ppm) | 249.939 | 234.9876 | 264.8904 | 242.6658 | 257.2123
W, Tungsten (ppm) | 94.4442 | 88.2906 | 100.5978 | 90.6859 | 98.2025
Y, Yttrium (ppm) | 34.6308 | 32.7882 | 36.4735 | 33.3923 | 35.8694
Y2O3, Yttrium(III) oxide (ppm) | 43.9777 | 41.6377 | 46.3177 | 42.4048 | 45.5505
Yb, Ytterbium (ppm) | 3.2944 | 2.9581 | 3.6306 | 3.0723 | 3.5164
Yb2O3, Ytterbium(III) oxide (ppm) | 3.7168 | 3.301 | 4.1325 | 3.4663 | 3.9673
Zn, Zinc (ppm) | 47.4348 | 39.2178 | 55.6518 | 43.7282 | 51.1413
Zr, Zirconium (ppm) | 282.6915 | 265.3087 | 300.0744 | 269.4957 | 295.8874
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 15.8094 | 15.5958 | 16.0231 | 15.7052 | 15.9137
BaO, Barium oxide (ppm) | 754.9696 | 663.4328 | 846.5065 | IND | IND
CaO, Calcium oxide (wt.%) | 0.31843 | 0.29842 | 0.33843 | 0.30202 | 0.33483
CeO2, Cerium(IV) oxide (wt.%) | 0.04363 | 0.034037 | 0.053222 | IND | IND
Cr2O3, Chromium(III) oxide (ppm) | 197.0764 | 151.1365 | 243.0163 | IND | IND
Dy2O3, Dysprosium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Er2O3, Erbium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Eu2O3, Europium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 7.2873 | 7.1774 | 7.3971 | 7.2167 | 7.3579
Ho2O3, Holmium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 2.8724 | 2.8191 | 2.9257 | 2.8527 | 2.8921
La2O3, Lanthanum(III) oxide (wt.%) | 0.030001 | 0.029297 | 0.030705 | IND | IND
Lu2O3, Lutetium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
MgO, Magnesium oxide (wt.%) | 0.82357 | 0.80333 | 0.84381 | 0.81191 | 0.83524
MnO, Manganese oxide (wt.%) | 0.011486 | 0.0063252 | 0.016647 | IND | IND
Na2O, Sodium oxide (wt.%) | 0.23613 | 0.2011 | 0.27116 | 0.22406 | 0.24819
Nb2O5, Niobium(V) oxide (ppm) | 147.3543 | 117.1658 | 177.5429 | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.12647 | 0.11863 | 0.13431 | 0.12359 | 0.12936
SiO2, Silicon dioxide (wt.%) | 66.403 | 65.8871 | 66.9189 | 66.1555 | 66.6506
Sm2O3, Samarium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
SnO2, Tin dioxide (ppm) | < 100 | IND | IND | IND | IND
SrO, Strontium oxide (ppm) | 112.3179 | 83.3394 | 141.2965 | IND | IND
Tb4O7, Terbium(III,IV) oxide (ppm) | < 100 | IND | IND | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.98577 | 0.96717 | 1.0044 | 0.97439 | 0.99715
Tm2O3, Thulium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
U3O8, Uranium(V,VI) oxide (ppm) | < 100 | IND | IND | IND | IND
Yb2O3, Ytterbium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
ZrO2, Zirconium dioxide (ppm) | 363.2181 | 292.1253 | 434.3108 | 340.4613 | 385.9748
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 4.0004 | 3.8158 | 4.185 | 3.9262 | 4.0746
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 466
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Cd | ppm | 0.040178 | Re | ppm | 0.0018958
Hg | ppm | 0.46833 | Se | ppm | 0.90185
Borate / Peroxide Fusion ICP
Ag | ppm | < 1 | Na2O | wt.% | 0.266 | Sb | ppm | 5.7792
Ge | ppm | 2.3333 | Pb | ppm | 25.7208 | Se | ppm | 13.7217
In | ppm | 0.19508 | Re | ppm | 0.0039483 | Te | ppm | < 1
Na | wt.% | 0.19733 | S | wt.% | 0.010505 | Tl | ppm | 0.77083
Borate Fusion XRF
As2O3 | ppm | 600 | HfO2 | ppm | < 100 | SO3 | wt.% | 0.012625
Bi2O3 | ppm | < 100 | MoO3 | ppm | < 100 | Ta2O5 | ppm | < 100
Cl | ppm | < 100 | Nd2O3 | wt.% | 0.016942 | ThO2 | ppm | 55.9239
Co3O4 | ppm | < 100 | NiO | ppm | 38.5363 | TOT_XRF | wt.% | 99.335
Cs2O | ppm | < 100 | PbO | ppm | < 100 | V2O5 | ppm | 244.4444
CuO | ppm | 46.9995 | Pr6O11 | ppm | 111.7896 | WO3 | ppm | 120.8333
F | wt.% | < 0.5 | Rb2O | ppm | 133.3333 | Y2O3 | ppm | 83.1403
Ga2O3 | ppm | < 100 | Sb2O3 | ppm | < 100 | ZnO | ppm | 59.204
Gd2O3 | ppm | 109.8796 | Sc2O3 | ppm | < 100
Thermogravimetry
H2O- | wt.% | 0.64245
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 466
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 0.32234 | 0.040356 | 0.24163 | 0.40305 | 0.20127 | 0.44341 | 12.52% | 25.04% | 37.56% | 0.30622 | 0.33846
Al, wt.% | 7.8995 | 0.28794 | 7.3236 | 8.4754 | 7.0356 | 8.7633 | 3.645% | 7.29% | 10.94% | 7.5045 | 8.2944
As, ppm | 481.4941 | 17.6419 | 446.2103 | 516.7778 | 428.5685 | 534.4197 | 3.664% | 7.328% | 10.99% | 457.4194 | 505.5688
Ba, ppm | 701.7594 | 26.2005 | 649.3584 | 754.1604 | 623.1578 | 780.3609 | 3.734% | 7.467% | 11.2% | 666.6714 | 736.8474
Be, ppm | 2.9045 | 0.17678 | 2.5509 | 3.258 | 2.3741 | 3.4348 | 6.087% | 12.17% | 18.26% | 2.7592 | 3.0497
Bi, ppm | 22.0859 | 1.6637 | 18.7586 | 25.4133 | 17.095 | 27.0769 | 7.533% | 15.07% | 22.6% | 20.9816 | 23.1902
Ca, wt.% | 0.22587 | 0.0081298 | 0.20961 | 0.24213 | 0.20148 | 0.25026 | 3.599% | 7.199% | 10.8% | 0.21458 | 0.23717
Ce, ppm | 339.3862 | 19.5882 | 300.2098 | 378.5625 | 280.6217 | 398.1506 | 5.772% | 11.54% | 17.31% | 322.4169 | 356.3555
Co, ppm | 2.9229 | 0.12182 | 2.6793 | 3.1666 | 2.5575 | 3.2884 | 4.168% | 8.335% | 12.5% | 2.7768 | 3.0691
Cr, ppm | 131.7938 | 14.5698 | 102.6541 | 160.9334 | 88.0843 | 175.5033 | 11.06% | 22.11% | 33.17% | 125.2041 | 138.3835
Cs, ppm | 10.7543 | 0.57409 | 9.6062 | 11.9025 | 9.0321 | 12.4766 | 5.338% | 10.68% | 16.01% | 10.2166 | 11.2921
Cu, ppm | 34.1434 | 1.3861 | 31.3712 | 36.9157 | 29.985 | 38.3018 | 4.06% | 8.119% | 12.18% | 32.4363 | 35.8506
Dy, ppm | 5.5683 | 0.26043 | 5.0475 | 6.0892 | 4.787 | 6.3496 | 4.677% | 9.354% | 14.03% | 5.2899 | 5.8467
Er, ppm | 2.1932 | 0.32224 | 1.5487 | 2.8377 | 1.2264 | 3.1599 | 14.69% | 29.39% | 44.08% | 2.0835 | 2.3028
Eu, ppm | 4.2557 | 0.17099 | 3.9137 | 4.5977 | 3.7427 | 4.7687 | 4.018% | 8.036% | 12.05% | 4.0429 | 4.4685
Fe, wt.% | 4.9673 | 0.17071 | 4.6259 | 5.3087 | 4.4552 | 5.4794 | 3.437% | 6.873% | 10.31% | 4.7189 | 5.2157
Ga, ppm | 22.3564 | 1.3019 | 19.7526 | 24.9601 | 18.4507 | 26.262 | 5.823% | 11.65% | 17.47% | 21.2385 | 23.4742
Gd, ppm | 11.4013 | 0.87703 | 9.6472 | 13.1553 | 8.7702 | 14.0324 | 7.692% | 15.38% | 23.08% | 10.8312 | 11.9714
Ge, ppm | 0.42109 | 0.083826 | 0.25344 | 0.58874 | 0.16962 | 0.67257 | 19.91% | 39.81% | 59.72% | 0.40004 | 0.44215
Hf, ppm | 4.6995 | 0.27765 | 4.1442 | 5.2548 | 3.8666 | 5.5325 | 5.908% | 11.82% | 17.72% | 4.4646 | 4.9345
Ho, ppm | 0.86995 | 0.096783 | 0.67639 | 1.0635 | 0.5796 | 1.1603 | 11.13% | 22.25% | 33.38% | 0.82646 | 0.91345
In, ppm | 0.12086 | 0.011623 | 0.097616 | 0.14411 | 0.085993 | 0.15573 | 9.617% | 19.23% | 28.85% | 0.11482 | 0.1269
K, wt.% | 2.3934 | 0.085969 | 2.2215 | 2.5653 | 2.1355 | 2.6513 | 3.592% | 7.184% | 10.78% | 2.2737 | 2.5131
La, ppm | 229.0675 | 16.6638 | 195.7399 | 262.3951 | 179.0761 | 279.0589 | 7.275% | 14.55% | 21.82% | 217.6141 | 240.5208
Li, ppm | 43.8747 | 1.6497 | 40.5753 | 47.1741 | 38.9256 | 48.8238 | 3.76% | 7.52% | 11.28% | 41.681 | 46.0685
Lu, ppm | 0.28168 | 0.033717 | 0.21425 | 0.34912 | 0.18053 | 0.38284 | 11.97% | 23.94% | 35.91% | 0.2676 | 0.29577
Mg, wt.% | 0.41156 | 0.030294 | 0.35097 | 0.47215 | 0.32068 | 0.50244 | 7.361% | 14.72% | 22.08% | 0.39098 | 0.43214
Mn, wt.% | 0.0091496 | 0.00052166 | 0.0081063 | 0.010193 | 0.0075846 | 0.010715 | 5.701% | 11.4% | 17.1% | 0.0086921 | 0.0096071
Mo, ppm | 6.4168 | 0.35621 | 5.7043 | 7.1292 | 5.3481 | 7.4854 | 5.551% | 11.1% | 16.65% | 6.0959 | 6.7376
Na, wt.% | 0.1638 | 0.01336 | 0.13708 | 0.19052 | 0.12372 | 0.20388 | 8.156% | 16.31% | 24.47% | 0.15561 | 0.17199
Nb, ppm | 109.4805 | 8.5749 | 92.3306 | 126.6303 | 83.7557 | 135.2052 | 7.832% | 15.66% | 23.5% | 104.0064 | 114.9545
Nd, ppm | 152.0785 | 7.9236 | 136.2313 | 167.9257 | 128.3076 | 175.8494 | 5.21% | 10.42% | 15.63% | 144.4746 | 159.6824
Ni, ppm | 19.0014 | 1.08 | 16.8414 | 21.1614 | 15.7615 | 22.2414 | 5.684% | 11.37% | 17.05% | 18.0513 | 19.9515
P, wt.% | 0.054981 | 0.0024059 | 0.05017 | 0.059793 | 0.047764 | 0.062199 | 4.376% | 8.752% | 13.13% | 0.052232 | 0.057731
Pb, ppm | 22.988 | 2.1867 | 18.6146 | 27.3614 | 16.4279 | 29.5481 | 9.512% | 19.02% | 28.54% | 21.8386 | 24.1374
Pr, ppm | 44.1153 | 1.196 | 41.7233 | 46.5073 | 40.5273 | 47.7033 | 2.711% | 5.422% | 8.133% | 41.9095 | 46.321
Rb, ppm | 140.2327 | 5.9933 | 128.2461 | 152.2193 | 122.2528 | 158.2126 | 4.274% | 8.548% | 12.82% | 133.2211 | 147.2443
S, wt.% | 0.011701 | 0.003158 | 0.0053852 | 0.018017 | 0.0022273 | 0.021175 | 26.99% | 53.98% | 80.97% | 0.011116 | 0.012286
Sb, ppm | 5.4451 | 0.19819 | 5.0487 | 5.8414 | 4.8505 | 6.0396 | 3.64% | 7.28% | 10.92% | 5.1728 | 5.7173
Sc, ppm | 18.8402 | 1.3726 | 16.0951 | 21.5854 | 14.7225 | 22.958 | 7.285% | 14.57% | 21.86% | 17.8982 | 19.7822
Sm, ppm | 20.8147 | 0.75881 | 19.297 | 22.3323 | 18.5382 | 23.0911 | 3.646% | 7.291% | 10.94% | 19.7739 | 21.8554
Sn, ppm | 17.6454 | 0.92686 | 15.7917 | 19.4991 | 14.8648 | 20.426 | 5.253% | 10.51% | 15.76% | 16.7631 | 18.5277
Sr, ppm | 107.0192 | 5.002 | 97.0151 | 117.0232 | 92.0131 | 122.0252 | 4.674% | 9.348% | 14.02% | 101.6682 | 112.3701
Ta, ppm | 2.5637 | 0.27234 | 2.019 | 3.1083 | 1.7466 | 3.3807 | 10.62% | 21.25% | 31.87% | 2.4355 | 2.6918
Tb, ppm | 1.2646 | 0.10545 | 1.0537 | 1.4755 | 0.94822 | 1.5809 | 8.339% | 16.68% | 25.02% | 1.2013 | 1.3278
Te, ppm | 0.30302 | 0.052307 | 0.1984 | 0.40763 | 0.14609 | 0.45994 | 17.26% | 34.52% | 51.79% | 0.28787 | 0.31817
Th, ppm | 32.4156 | 2.0611 | 28.2934 | 36.5378 | 26.2324 | 38.5989 | 6.358% | 12.72% | 19.07% | 30.7948 | 34.0364
Ti, wt.% | 0.44801 | 0.042732 | 0.36255 | 0.53348 | 0.31982 | 0.57621 | 9.538% | 19.08% | 28.61% | 0.42561 | 0.47041
Tl, ppm | 0.78802 | 0.059279 | 0.66946 | 0.90658 | 0.61018 | 0.96585 | 7.523% | 15.05% | 22.57% | 0.74862 | 0.82742
Tm, ppm | 0.27431 | 0.028308 | 0.21769 | 0.33092 | 0.18938 | 0.35923 | 10.32% | 20.64% | 30.96% | 0.26059 | 0.28802
U, ppm | 3.0374 | 0.22003 | 2.5973 | 3.4775 | 2.3773 | 3.6975 | 7.244% | 14.49% | 21.73% | 2.8855 | 3.1893
V, ppm | 127.5664 | 3.9903 | 119.5859 | 135.547 | 115.5956 | 139.5373 | 3.128% | 6.256% | 9.384% | 121.1881 | 133.9448
W, ppm | 90.1577 | 6.7055 | 76.7468 | 103.5687 | 70.0413 | 110.2741 | 7.437% | 14.87% | 22.31% | 85.6498 | 94.6656
Y, ppm | 20.2983 | 1.3761 | 17.546 | 23.0506 | 16.1699 | 24.4267 | 6.78% | 13.56% | 20.34% | 19.2834 | 21.3132
Yb, ppm | 1.8252 | 0.16098 | 1.5032 | 2.1471 | 1.3423 | 2.3081 | 8.82% | 17.64% | 26.46% | 1.7339 | 1.9165
Zn, ppm | 41.8345 | 2.1852 | 37.464 | 46.2049 | 35.2788 | 48.3901 | 5.223% | 10.45% | 15.67% | 39.7427 | 43.9262
Zr, ppm | 161.3336 | 7.5193 | 146.295 | 176.3721 | 138.7758 | 183.8914 | 4.661% | 9.321% | 13.98% | 153.2669 | 169.4003
Borate / Peroxide Fusion ICP
Al, wt.% | 8.345 | 0.17147 | 8.002 | 8.6879 | 7.8305 | 8.8594 | 2.055% | 4.11% | 6.164% | 7.9277 | 8.7622
Al2O3, wt.% | 15.7678 | 0.32399 | 15.1198 | 16.4158 | 14.7958 | 16.7398 | 2.055% | 4.11% | 6.164% | 14.9794 | 16.5562
As, ppm | 472.25 | 34.4445 | 403.3611 | 541.1389 | 368.9166 | 575.5834 | 7.294% | 14.59% | 21.88% | 448.6375 | 495.8625
B, wt.% | 0.1383 | 0.01121 | 0.11588 | 0.16072 | 0.10467 | 0.17193 | 8.105% | 16.21% | 24.32% | 0.13139 | 0.14522
Ba, ppm | 718.8531 | 16.0957 | 686.6617 | 751.0446 | 670.566 | 767.1403 | 2.239% | 4.478% | 6.717% | 682.9105 | 754.7958
BaO, ppm | 802.5995 | 17.9709 | 766.6578 | 838.5413 | 748.6869 | 856.5121 | 2.239% | 4.478% | 6.717% | 762.4696 | 842.7295
Be, ppm | 3.1236 | 0.38351 | 2.3566 | 3.8906 | 1.9731 | 4.2741 | 12.28% | 24.56% | 36.83% | 2.9674 | 3.2798
Bi, ppm | 21.9567 | 1.0611 | 19.8344 | 24.0789 | 18.7733 | 25.14 | 4.833% | 9.666% | 14.5% | 20.8588 | 23.0545
Ca, wt.% | 0.24507 | 0.017308 | 0.21046 | 0.27969 | 0.19315 | 0.297 | 7.062% | 14.12% | 21.19% | 0.23282 | 0.25733
CaO, wt.% | 0.34291 | 0.024217 | 0.29447 | 0.39134 | 0.27026 | 0.41556 | 7.062% | 14.12% | 21.19% | 0.32576 | 0.36005
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, wt.% | 0.035574 | 0.0024432 | 0.030687 | 0.04046 | 0.028244 | 0.042903 | 6.868% | 13.74% | 20.6% | 0.033795 | 0.037353
CeO2, wt.% | 0.043699 | 0.0030012 | 0.037696 | 0.049701 | 0.034695 | 0.052703 | 6.868% | 13.74% | 20.6% | 0.041514 | 0.045884
Co, ppm | 3.2566 | 0.37964 | 2.4973 | 4.0159 | 2.1177 | 4.3955 | 11.66% | 23.32% | 34.97% | 3.0938 | 3.4194
Cr, ppm | 170.2198 | 23.0033 | 124.2133 | 216.2263 | 101.21 | 239.2296 | 13.51% | 27.03% | 40.54% | 161.7088 | 178.7308
Cr2O3, ppm | 248.7762 | 33.6193 | 181.5377 | 316.0147 | 147.9184 | 349.634 | 13.51% | 27.03% | 40.54% | 236.3374 | 261.215
Cs, ppm | 10.8718 | 0.34008 | 10.1916 | 11.552 | 9.8516 | 11.8921 | 3.128% | 6.256% | 9.384% | 10.3282 | 11.4154
Cu, ppm | 35.5507 | 3.075 | 29.4006 | 41.7007 | 26.3256 | 44.7757 | 8.65% | 17.3% | 25.95% | 33.7731 | 37.3282
Dy, ppm | 7.6412 | 0.43516 | 6.7709 | 8.5115 | 6.3357 | 8.9467 | 5.695% | 11.39% | 17.08% | 7.2592 | 8.0233
Dy2O3, ppm | 8.7208 | 0.51643 | 7.6879 | 9.7536 | 7.1715 | 10.27 | 5.922% | 11.84% | 17.77% | 8.2847 | 9.1568
Er, ppm | 3.6537 | 0.20088 | 3.2519 | 4.0554 | 3.051 | 4.2563 | 5.498% | 11% | 16.49% | 3.471 | 3.8364
Er2O3, ppm | 4.2414 | 0.29854 | 3.6443 | 4.8384 | 3.3458 | 5.137 | 7.039% | 14.08% | 21.12% | 4.0293 | 4.4534
Eu, ppm | 4.4164 | 0.2015 | 4.0134 | 4.8194 | 3.8119 | 5.0209 | 4.563% | 9.125% | 13.69% | 4.1956 | 4.6372
Eu2O3, ppm | 5.1137 | 0.23332 | 4.6471 | 5.5804 | 4.4138 | 5.8137 | 4.563% | 9.125% | 13.69% | 4.858 | 5.3694
Fe, wt.% | 5.1729 | 0.14049 | 4.892 | 5.4539 | 4.7515 | 5.5944 | 2.716% | 5.432% | 8.147% | 4.9143 | 5.4316
Fe2O3, wt.% | 7.3957 | 0.20085 | 6.994 | 7.7974 | 6.7932 | 7.9983 | 2.716% | 5.432% | 8.147% | 7.0259 | 7.7655
Ga, ppm | 22.2292 | 1.1627 | 19.9038 | 24.5546 | 18.7411 | 25.7173 | 5.23% | 10.46% | 15.69% | 21.1178 | 23.3407
Gd, ppm | 12.1392 | 0.49378 | 11.1517 | 13.1268 | 10.6579 | 13.6206 | 4.068% | 8.135% | 12.2% | 11.5323 | 12.7462
Gd2O3, ppm | 13.9743 | 0.52841 | 12.9175 | 15.0311 | 12.3891 | 15.5595 | 3.781% | 7.563% | 11.34% | 13.2756 | 14.673
Hf, ppm | 7.5001 | 0.43639 | 6.6273 | 8.3729 | 6.1909 | 8.8093 | 5.818% | 11.64% | 17.46% | 7.1251 | 7.8751
HfO2, ppm | 8.8449 | 0.51464 | 7.8156 | 9.8741 | 7.301 | 10.3888 | 5.818% | 11.64% | 17.46% | 8.4026 | 9.2871
Ho, ppm | 1.3651 | 0.066211 | 1.2327 | 1.4976 | 1.1665 | 1.5638 | 4.85% | 9.7% | 14.55% | 1.2969 | 1.4334
Ho2O3, ppm | 1.5638 | 0.075844 | 1.4121 | 1.7155 | 1.3362 | 1.7913 | 4.85% | 9.7% | 14.55% | 1.4856 | 1.642
K, wt.% | 2.4058 | 0.064392 | 2.277 | 2.5346 | 2.2126 | 2.599 | 2.677% | 5.353% | 8.03% | 2.2855 | 2.5261
K2O, wt.% | 2.898 | 0.077567 | 2.7429 | 3.0532 | 2.6653 | 3.1307 | 2.677% | 5.353% | 8.03% | 2.7531 | 3.0429
La, ppm | 237.9626 | 14.5807 | 208.8012 | 267.1241 | 194.2204 | 281.7048 | 6.127% | 12.25% | 18.38% | 226.0645 | 249.8608
La2O3, ppm | 279.0826 | 17.1003 | 244.882 | 313.2831 | 227.7817 | 330.3834 | 6.127% | 12.25% | 18.38% | 265.1284 | 293.0367
Li, ppm | 44.396 | 2.34 | 39.716 | 49.0761 | 37.376 | 51.4161 | 5.271% | 10.54% | 15.81% | 42.1762 | 46.6158
Lu, ppm | 0.46916 | 0.03585 | 0.39746 | 0.54086 | 0.36161 | 0.57671 | 7.641% | 15.28% | 22.92% | 0.4457 | 0.49262
Lu2O3, ppm | 0.53004 | 0.048565 | 0.43291 | 0.62717 | 0.38434 | 0.67573 | 9.163% | 18.33% | 27.49% | 0.50354 | 0.55654
Mg, wt.% | 0.48556 | 0.024246 | 0.43707 | 0.53405 | 0.41282 | 0.5583 | 4.993% | 9.987% | 14.98% | 0.46128 | 0.50984
MgO, wt.% | 0.80515 | 0.040205 | 0.72474 | 0.88556 | 0.68454 | 0.92577 | 4.993% | 9.987% | 14.98% | 0.76489 | 0.84541
Mn, wt.% | 0.010021 | 0.00035259 | 0.0093155 | 0.010726 | 0.0089629 | 0.011078 | 3.519% | 7.037% | 10.56% | 0.0095197 | 0.010522
MnO, wt.% | 0.012939 | 0.00045528 | 0.012028 | 0.01385 | 0.011573 | 0.014305 | 3.519% | 7.037% | 10.56% | 0.012292 | 0.013586
Mo, ppm | 6.7233 | 0.53799 | 5.6473 | 7.7993 | 5.1094 | 8.3373 | 8.002% | 16% | 24.01% | 6.3872 | 7.0595
Nb, ppm | 112.9384 | 5.9499 | 101.0386 | 124.8382 | 95.0887 | 130.7881 | 5.268% | 10.54% | 15.8% | 107.2915 | 118.5853
Nd, ppm | 155.156 | 9.5629 | 136.0301 | 174.2819 | 126.4672 | 183.8448 | 6.163% | 12.33% | 18.49% | 147.3982 | 162.9138
Nd2O3, ppm | 180.9739 | 11.1542 | 158.6655 | 203.2824 | 147.5113 | 214.4366 | 6.163% | 12.33% | 18.49% | 171.9252 | 190.0226
Ni, ppm | 24.6333 | 6.4299 | 11.7735 | 37.4931 | 5.3436 | 43.923 | 26.1% | 52.2% | 78.31% | 23.4017 | 25.865
P, wt.% | 0.056036 | 0.004545 | 0.046946 | 0.065126 | 0.042401 | 0.069671 | 8.111% | 16.22% | 24.33% | 0.053234 | 0.058838
P2O5, wt.% | 0.12841 | 0.010415 | 0.10758 | 0.14924 | 0.097166 | 0.15966 | 8.111% | 16.22% | 24.33% | 0.12199 | 0.13483
Pr, ppm | 44.3622 | 2.5048 | 39.3526 | 49.3719 | 36.8478 | 51.8767 | 5.646% | 11.29% | 16.94% | 42.1441 | 46.5803
Pr6O11, ppm | 53.5984 | 3.0263 | 47.5458 | 59.6511 | 44.5195 | 62.6774 | 5.646% | 11.29% | 16.94% | 50.9185 | 56.2784
Rb, ppm | 137.7093 | 6.878 | 123.9533 | 151.4653 | 117.0754 | 158.3433 | 4.995% | 9.989% | 14.98% | 130.8239 | 144.5948
Sc, ppm | 19.1837 | 3.2132 | 12.7574 | 25.61 | 9.5442 | 28.8232 | 16.75% | 33.5% | 50.25% | 18.2245 | 20.1429
Si, wt.% | 30.6954 | 1.609 | 27.4775 | 33.9134 | 25.8685 | 35.5224 | 5.242% | 10.48% | 15.73% | 29.1606 | 32.2302
SiO2, wt.% | 65.6636 | 3.4419 | 58.7798 | 72.5475 | 55.3378 | 75.9895 | 5.242% | 10.48% | 15.73% | 62.3805 | 68.9468
Sm, ppm | 20.9838 | 1.2875 | 18.4088 | 23.5589 | 17.1212 | 24.8464 | 6.136% | 12.27% | 18.41% | 19.9346 | 22.033
Sm2O3, ppm | 24.3329 | 1.493 | 21.3468 | 27.3189 | 19.8538 | 28.8119 | 6.136% | 12.27% | 18.41% | 23.1162 | 25.5495
Sn, ppm | 20.4984 | 1.3172 | 17.8639 | 23.1329 | 16.5466 | 24.4501 | 6.426% | 12.85% | 19.28% | 19.4735 | 21.5233
Sr, ppm | 109.8755 | 5.9305 | 98.0146 | 121.7365 | 92.0841 | 127.6669 | 5.397% | 10.79% | 16.19% | 104.3818 | 115.3693
Ta, ppm | 2.9421 | 0.17266 | 2.5967 | 3.2874 | 2.4241 | 3.4601 | 5.869% | 11.74% | 17.61% | 2.795 | 3.0892
Tb, ppm | 1.4826 | 0.090802 | 1.301 | 1.6642 | 1.2102 | 1.755 | 6.125% | 12.25% | 18.37% | 1.4084 | 1.5567
Tb4O7, ppm | 1.7438 | 0.1068 | 1.5302 | 1.9574 | 1.4234 | 2.0642 | 6.125% | 12.25% | 18.37% | 1.6566 | 1.831
Th, ppm | 31.9164 | 1.6507 | 28.6151 | 35.2178 | 26.9644 | 36.8685 | 5.172% | 10.34% | 15.52% | 30.3206 | 33.5123
Ti, wt.% | 0.58653 | 0.01577 | 0.55499 | 0.61807 | 0.53922 | 0.63384 | 2.689% | 5.378% | 8.066% | 0.5572 | 0.61585
TiO2, wt.% | 0.97839 | 0.026307 | 0.92577 | 1.031 | 0.89947 | 1.0573 | 2.689% | 5.378% | 8.066% | 0.92947 | 1.0273
Tm, ppm | 0.50597 | 0.029907 | 0.44615 | 0.56578 | 0.41625 | 0.59569 | 5.911% | 11.82% | 17.73% | 0.48067 | 0.53127
Tm2O3, ppm | 0.57612 | 0.03541 | 0.5053 | 0.64694 | 0.46989 | 0.68235 | 6.146% | 12.29% | 18.44% | 0.54732 | 0.60493
U, ppm | 3.4449 | 0.22272 | 2.9995 | 3.8903 | 2.7768 | 4.1131 | 6.465% | 12.93% | 19.4% | 3.2727 | 3.6172
V, ppm | 140.0062 | 5.0375 | 129.9312 | 150.0812 | 124.8937 | 155.1186 | 3.598% | 7.196% | 10.79% | 133.0059 | 147.0065
V2O5, ppm | 249.939 | 8.9929 | 231.9532 | 267.9249 | 222.9602 | 276.9178 | 3.598% | 7.196% | 10.79% | 237.4421 | 262.436
W, ppm | 94.4442 | 5.0617 | 84.3208 | 104.5676 | 79.2591 | 109.6293 | 5.359% | 10.72% | 16.08% | 89.722 | 99.1664
Y, ppm | 34.6308 | 1.9831 | 30.6647 | 38.5969 | 28.6817 | 40.58 | 5.726% | 11.45% | 17.18% | 32.8993 | 36.3624
Y2O3, ppm | 43.9777 | 2.5183 | 38.9411 | 49.0143 | 36.4229 | 51.5325 | 5.726% | 11.45% | 17.18% | 41.7788 | 46.1766
Yb, ppm | 3.2944 | 0.17588 | 2.9426 | 3.6461 | 2.7667 | 3.822 | 5.339% | 10.68% | 16.02% | 3.1296 | 3.4591
Yb2O3, ppm | 3.7168 | 0.25633 | 3.2041 | 4.2294 | 2.9478 | 4.4858 | 6.897% | 13.79% | 20.69% | 3.5309 | 3.9026
Zn, ppm | 47.4348 | 4.0218 | 39.3912 | 55.4783 | 35.3695 | 59.5001 | 8.479% | 16.96% | 25.44% | 45.063 | 49.8065
Zr, ppm | 282.6915 | 12.402 | 257.8875 | 307.4956 | 245.4855 | 319.8976 | 4.387% | 8.774% | 13.16% | 268.557 | 296.8261
Borate Fusion XRF
Al2O3, wt.% | 15.8094 | 0.25298 | 15.3035 | 16.3154 | 15.0505 | 16.5684 | 1.6% | 3.2% | 4.801% | 15.019 | 16.5999
BaO, ppm | 754.9696 | 70.3473 | 614.2751 | 895.6642 | 543.9278 | 966.0115 | 9.318% | 18.64% | 27.95% | 717.2212 | 792.7181
CaO, wt.% | 0.31843 | 0.019576 | 0.27927 | 0.35758 | 0.2597 | 0.37715 | 6.148% | 12.3% | 18.44% | 0.3025 | 0.33435
CeO2, wt.% | 0.04363 | 0.0074696 | 0.02869 | 0.058569 | 0.021221 | 0.066039 | 17.12% | 34.24% | 51.36% | 0.041448 | 0.045811
Cr2O3, ppm | 197.0764 | 50.1191 | 96.8382 | 297.3146 | 46.7192 | 347.4337 | 25.43% | 50.86% | 76.29% | 187.2226 | 206.9302
Dy2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Er2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Eu2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Fe2O3, wt.% | 7.2873 | 0.099905 | 7.0875 | 7.4871 | 6.9876 | 7.587 | 1.371% | 2.742% | 4.113% | 6.9229 | 7.6516
Ho2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 2.8724 | 0.059804 | 2.7528 | 2.992 | 2.693 | 3.0518 | 2.082% | 4.164% | 6.246% | 2.7288 | 3.016
La2O3, wt.% | 0.030001 | 0.00037676 | 0.029247 | 0.030754 | 0.028871 | 0.031131 | 1.256% | 2.512% | 3.768% | 0.028501 | 0.031501
Lu2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
MgO, wt.% | 0.82357 | 0.021615 | 0.78034 | 0.8668 | 0.75873 | 0.88842 | 2.624% | 5.249% | 7.873% | 0.78239 | 0.86475
MnO, wt.% | 0.011486 | 0.0028638 | 0.0057583 | 0.017213 | 0.0028946 | 0.020077 | 24.93% | 49.87% | 74.8% | 0.010912 | 0.01206
Na2O, wt.% | 0.23613 | 0.031441 | 0.17324 | 0.29901 | 0.1418 | 0.33045 | 13.32% | 26.63% | 39.95% | 0.22432 | 0.24793
Nb2O5, ppm | 147.3543 | 32.1638 | 83.0266 | 211.682 | 50.8628 | 243.8459 | 21.83% | 43.66% | 65.48% | 139.9866 | 154.7221
P2O5, wt.% | 0.12647 | 0.0054144 | 0.11564 | 0.1373 | 0.11023 | 0.14271 | 4.281% | 8.562% | 12.84% | 0.12015 | 0.13279
SiO2, wt.% | 66.403 | 0.51414 | 65.3747 | 67.4313 | 64.8606 | 67.9454 | 0.7743% | 1.549% | 2.323% | 63.0829 | 69.7232
Sm2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SnO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SrO, ppm | 112.3179 | 22.1112 | 68.0956 | 156.5403 | 45.9844 | 178.6515 | 19.69% | 39.37% | 59.06% | 106.702 | 117.9338
Tb4O7, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
TiO2, wt.% | 0.98577 | 0.015285 | 0.9552 | 1.0163 | 0.93992 | 1.0316 | 1.551% | 3.101% | 4.652% | 0.93648 | 1.0351
Tm2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
U3O8, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Yb2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
ZrO2, ppm | 363.2181 | 39.0892 | 285.0396 | 441.3965 | 245.9504 | 480.4857 | 10.76% | 21.52% | 32.29% | 345.0572 | 381.379
Thermogravimetry
LOI1000, wt.% | 4.0004 | 0.30733 | 3.3857 | 4.6151 | 3.0784 | 4.9224 | 7.682% | 15.36% | 23.05% | 3.8004 | 4.2004
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 466
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Text Values:
NA | Unable to report due to QC failure (Lab 2.02)
NA | Not Applicable (Lab 2.16)
NA | Not Applicable (Lab 2.17)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 466
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Inspectorate (BV), Lima, Peru
Intertek, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Ankara, Anatolia, Turkey
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 466 (Certified Value 0.322 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | 0.35 | 0.4 | 0.37951 | 0.5 | 0.336 | 0.26 | 0.33 | 0.29 | 0.25 | 0.517 | 0.3 | 0.3 | 0.31 | 0.97 | 0.36 | [5 empty] | 1
| 1 | 2 | 0.35 | 0.4 | 0.37842 | 0.3 | 0.317 | 0.26 | 0.34 | 0.32 | 0.28 | 0.567 | 0.3 | 0.3 | 0.32 | 0.99 | 0.64 | [5 empty] | 4
| 1 | 3 | 0.34 | 0.4 | 0.3765 | 0.6 | 0.328 | 0.27 | 0.36 | 0.24 | 0.28 | 0.545 | 0.3 | 0.3 | 0.33 | 1.35 | 0.33 | [5 empty] | 16
| 1 | 4 | 0.34 | 0.7 | 0.37809 | 0.6 | 0.343 | 0.37 | 0.23 | 0.29 | 0.2 | 0.53 | 0.3 | 0.3 | 0.3 | 0.98 | 0.34 | [5 empty] | 0.32234
| 1 | 5 | 0.34 | 0.4 | 0.36226 | 0.6 | 0.315 | 0.28 | 0.29 | 0.31 | 0.27 | 0.573 | 0.3 | 0.3 | 0.35 | 1.48 | 0.36 | [5 empty] | 7
| 1 | 6 | 0.34 | 0.4 | 0.36752 | 0.5 | 0.355 | 0.18 | 0.36 | 0.29 | 0.29 | 0.529 | 0.3 | 0.3 | 0.33 | 1.03 | 0.4
| Mean |  | 0.34333 | 0.45 | 0.37372 | 0.51667 | 0.33233 | 0.27 | 0.31833 | 0.29 | 0.26167 | 0.5435 | 0.3 | 0.3 | 0.32333 | 1.1333 | 0.405
| Median |  | 0.34 | 0.4 | 0.3773 | 0.55 | 0.332 | 0.265 | 0.335 | 0.29 | 0.275 | 0.5375 | 0.3 | 0.3 | 0.325 | 1.01 | 0.36
| Std Dev. |  | 0.005164 | 0.12247 | 0.0071041 | 0.1169 | 0.015462 | 0.060663 | 0.050365 | 0.027568 | 0.033116 | 0.022448 | 0 | 0 | 0.017512 | 0.22295 | 0.1176
| Rel.Std.Dev. |  | 1.504% | 27.22% | 1.901% | 22.63% | 4.652% | 22.47% | 15.82% | 9.506% | 12.66% | 4.13% | 0% | 0% | 5.416% | 19.67% | 29.04%
| PDM3 |  | 6.513% | 39.6% | 15.94% | 60.29% | 3.101% | -16.24% | -1.243% | -10.03% | -18.82% | 68.61% | -6.93% | -6.93% | 0.3085% | 251.6% | 25.64%
| Z-Score (Absolute) |  | 0.18 | NA | 0.66 | NA | 0 | 0.99 | 0.22 | 0.67 | 1.13 | 3.36 | NA | NA | 0.14 | 12.76 | 1.16
Results from laboratories 2, 7, 16 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 466 (Certified Value 7.9 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 7.9591 | 8.2206 | 8.0275 | 7.5633 | 8.3321 | 8.13 | 7.41 | 7.66 | 8.12 | 7.85 | 8.08 | 9.34 | 7.16 | 7.79 | 7.37 | 8.199 | 7.95 | 7.91 |  |  | 1
| 1 | 2 | 8.0329 | 8.3183 | 8.0319 | 7.5719 | 8.0936 | 8.11 | 7.43 | 7.61 | 7.72 | 7.79 | 7.9 | 8.52 | 6.9 | 7.79 | 7.96 | 7.9389 | 7.97 | 7.24 |  |  | 0
| 1 | 3 | 7.9821 | 8.3888 | 8.024 | 7.5242 | 8.2416 | 7.94 | 7.3 | 7.55 | 7.88 | 7.67 | 8.03 | 8.72 | 7.04 | 7.84 | 8.44 | 7.9109 | 7.91 | 7.57 |  |  | 16
| 1 | 4 | 7.8833 | 8.0663 | 8.0262 | 7.5341 | 8.2521 | 8.15 | 7.33 | 7.54 | 8.73 | 7.94 | 7.9 | 8.68 | 6.92 | 7.77 | 7.6 | 8.39 | 7.93 | 8 |  |  | 7.8995
| 1 | 5 | 7.9755 | 8.3518 | 8.0256 | 7.4418 | 8.2672 | 8.04 | 7.35 | 7.66 | 8.14 | 7.83 | 8.09 | 8.61 | 7.24 | 7.68 | 7.76 | 8.1275 | 7.98 | 7.77 |  |  | 8
| 1 | 6 | 7.9305 | 8.0929 | 8.0261 | 7.3466 | 8.2319 | 7.96 | 7.55 | 7.59 | 8.13 | 7.79 | 8.02 | 8.68 | 6.94 | 7.78 | 7.89 | 8.0491 | 7.97 | 8.18
| Mean |  | 7.9606 | 8.2398 | 8.0269 | 7.497 | 8.2364 | 8.055 | 7.395 | 7.6017 | 8.12 | 7.8117 | 8.0033 | 8.7583 | 7.0333 | 7.775 | 7.8367 | 8.1026 | 7.9517 | 7.7783
| Median |  | 7.9673 | 8.2695 | 8.0261 | 7.5292 | 8.2468 | 8.075 | 7.38 | 7.6 | 8.125 | 7.81 | 8.025 | 8.68 | 6.99 | 7.785 | 7.825 | 8.0883 | 7.96 | 7.84
| Std Dev. |  | 0.050587 | 0.13635 | 0.0027074 | 0.086946 | 0.07847 | 0.08961 | 0.090277 | 0.051929 | 0.34357 | 0.088638 | 0.084538 | 0.29356 | 0.1401 | 0.05244 | 0.36379 | 0.17828 | 0.027142 | 0.33499
| Rel.Std.Dev. |  | 0.6355% | 1.655% | 0.03373% | 1.16% | 0.9527% | 1.112% | 1.221% | 0.6831% | 4.231% | 1.135% | 1.056% | 3.352% | 1.992% | 0.6745% | 4.642% | 2.2% | 0.3413% | 4.307%
| PDM3 |  | 0.7734% | 4.308% | 1.613% | -5.095% | 4.266% | 1.969% | -6.386% | -3.77% | 2.792% | -1.112% | 1.315% | 10.87% | -10.96% | -1.576% | -0.7951% | 2.571% | 0.6607% | -1.534%
| Z-Score (Absolute) |  | 0.02 | 1.12 | 0.28 | 1.81 | 1.11 | 0.39 | 2.21 | 1.4 | 0.65 | 0.57 | 0.19 | 3.17 | 3.64 | 0.71 | 0.47 | 0.58 | 0.02 | 0.7
| Analytical results for As in OREAS 466 (Certified Value 481 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 485 | 490.7 | 465.8136 | 362.56 | 470.0808 | 395 | 490.5 | 465 | 517 | 467 | 487 | 521.9 | 471 | 479 | 483 | 673 |  |  |  |  | 1
| 1 | 2 | 469 | 499.6 | 456.3864 | 360.15 | 471.891 | 361 | 493.5 | 467 | 488 | 465 | 474 | 528.1 | 475 | 477 | 484 | 712 |  |  |  |  | 0
| 1 | 3 | 474 | 497.5 | 447.5424 | 363.56 | 472.0969 | 414 | 492.6 | 464 | 501 | 453 | 482 | 499.6 | 468 | 487.5 | 483 | 671 |  |  |  |  | 16
| 1 | 4 | 468 | 484.1 | 451.9368 | 358.79 | 472.871 | 430 | 491.2 | 461 | 524 | 470 | 475 | 504.6 | 481 | 484.7 | 483 | 689 |  |  |  |  | 481.4941
| 1 | 5 | 494 | 491.9 | 473.3472 | 352.55 | 461.7907 | 408 | 489.9 | 468 | 515 | 465 | 485 | 517.7 | 472 | 488.2 | 482 | 689 |  |  |  |  | 9
| 1 | 6 | 475 | 490.4 | 440.8896 | 349.92 | 470.9918 | 349 | 496.9 | 462 | 513 | 465 | 481 | 495.4 | 479 | 486.4 | 482 | 679
| Mean |  | 477.5 | 492.3667 | 455.986 | 357.9217 | 469.9537 | 392.8333 | 492.4333 | 464.5 | 509.6667 | 464.1667 | 480.6667 | 511.2167 | 474.3333 | 483.8 | 482.8333 | 685.5
| Median |  | 474.5 | 491.3 | 454.1616 | 359.47 | 471.4414 | 401.5 | 491.9 | 464.5 | 514 | 465 | 481.5 | 511.15 | 473.5 | 485.55 | 483 | 684
| Std Dev. |  | 10.0946 | 5.5475 | 11.9478 | 5.5119 | 4.113 | 31.6191 | 2.561 | 2.7386 | 12.9872 | 5.8109 | 5.2409 | 13.1917 | 4.9666 | 4.6883 | 0.75277 | 15.0698
| Rel.Std.Dev. |  | 2.114% | 1.127% | 2.62% | 1.54% | 0.8752% | 8.049% | 0.5201% | 0.5896% | 2.548% | 1.252% | 1.09% | 2.58% | 1.047% | 0.9691% | 0.1559% | 2.198%
| PDM3 |  | -0.8295% | 2.258% | -5.298% | -25.66% | -2.397% | -18.41% | 2.272% | -3.529% | 5.851% | -3.599% | -0.1718% | 6.173% | -1.487% | 0.4789% | 0.2781% | 42.37%
| Z-Score (Absolute) |  | 0.08 | 0.64 | 1.12 | 5.85 | 0.44 | 4.16 | 0.64 | 0.7 | 1.48 | 0.72 | 0.08 | 1.55 | 0.23 | 0.23 | 0.18 | 9.97
| Analytical results for Ba in OREAS 466 (Certified Value 702 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 711.7 | 779 | 718.4248 | 563.9911 | 699.6144 | 741 | 710 | 670 | 720 | 690 | 710 | 741 | 690 | 723 | 658.7513 | 1325 | 685.4 | 694 | 720 |  | 1
| 1 | 2 | 729 | 787 | 715.4051 | 575.7168 | 693.5241 | 698 | 708 | 670 | 680 | 690 | 700 | 750 | 693 | 733 | 624.5665 | 1433 | 693.7 | 699 | 800 |  | 0
| 1 | 3 | 722.5 | 795 | 716.9456 | 545.0318 | 702.1322 | 662 | 712 | 660 | 700 | 680 | 710 | 714 | 688 | 728 | 628.0419 | 1371 | 692.7 | 697 | 710 |  | 16
| 1 | 4 | 723.2 | 772 | 718.3126 | 555.5985 | 693.7498 | 679 | 706 | 660 | 780 | 700 | 700 | 719 | 685 | 734 | 629.0045 | 1104 | 683.5 | 691 | 750 |  | 701.7594
| 1 | 5 | 715.6 | 789 | 716.4219 | 713.2915 | 686.0486 | 713 | 701 | 670 | 730 | 690 | 710 | 731 | 690 | 722 | 662.34 | 1387 | 697.5 | 700 | 740 |  | 10
| 1 | 6 | 708.6 | 781 | 714.763 | 494.1167 | 700.1201 | 679 | 738 | 670 | 720 | 690 | 710 | 705 | 690 | 713 | 666.635 | 1372 | 684.7 | 697 | 740
| Mean |  | 718.4333 | 783.8333 | 716.7122 | 574.6244 | 695.8649 | 695.3333 | 712.5 | 666.6667 | 721.6667 | 690 | 706.6667 | 726.6667 | 689.3333 | 725.5 | 644.8899 | 1332 | 689.5833 | 696.3333 | 743.3333
| Median |  | 719.05 | 784 | 716.6838 | 559.7948 | 696.6821 | 688.5 | 709 | 670 | 720 | 690 | 710 | 725 | 690 | 725.5 | 643.8779 | 1371.5 | 689.05 | 697 | 740
| Std Dev. |  | 7.7585 | 8.1588 | 1.4927 | 73.5691 | 5.9625 | 28.4582 | 13.0499 | 5.164 | 33.7145 | 6.3246 | 5.164 | 17.0724 | 2.6583 | 7.8677 | 19.5895 | 116.9444 | 5.7912 | 3.3267 | 31.4113
| Rel.Std.Dev. |  | 1.08% | 1.041% | 0.2083% | 12.8% | 0.8568% | 4.093% | 1.832% | 0.7746% | 4.672% | 0.9166% | 0.7308% | 2.349% | 0.3856% | 1.084% | 3.038% | 8.78% | 0.8398% | 0.4777% | 4.226%
| PDM3 |  | 2.376% | 11.7% | 2.131% | -18.12% | -0.84% | -0.9157% | 1.531% | -5.001% | 2.837% | -1.676% | 0.6993% | 3.549% | -1.771% | 3.383% | -8.104% | 89.81% | -1.735% | -0.7732% | 5.924%
| Z-Score (Absolute) |  | 0.46 | 3.05 | 0.4 | 5.21 | 0.43 | 0.45 | 0.23 | 1.58 | 0.59 | 0.66 | 0 | 0.79 | 0.68 | 0.74 | 2.44 | 24.68 | 0.67 | 0.41 | 1.45
| Analytical results for Be in OREAS 466 (Certified Value 2.9 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 3.36 | 2.65 | 3.0734 | 2.7523 | 2.8406 | 4 | 3 | 2.69 | 3 | 2.86 | 2.92 | 3 | 3 | 2.8 | 2.9424 | 3.5 | 2.98 | 3 | 3.56 |  | 1
| 1 | 2 | 3 | 2.67 | 3.037 | 2.6908 | 2.8918 | 4.1 | 3 | 2.69 | 2.87 | 2.83 | 2.85 | 3 | 3 | 2.9 | 3.0047 | 3.44 | 3.15 | 3 | 3.36 |  | 0
| 1 | 3 | 3 | 2.67 | 3.0401 | 2.8753 | 2.8395 | 3.7 | 3 | 2.66 | 2.92 | 2.8 | 2.89 | 3 | 3 | 2.9 | 3.0156 | 3.25 | 3.02 | 3 | 2.37 |  | 16
| 1 | 4 | 3.05 | 2.59 | 3.0612 | 3.275 | 2.9259 | 3.8 | 3 | 2.65 | 3.16 | 2.89 | 2.88 | 3 | 3 | 2.9 | 3.0114 | 3.43 | 2.9 | 3 | 2.86 |  | 2.9045
| 1 | 5 | 2.96 | 2.66 | 3.0001 | 3.152 | 2.9444 | 3.9 | 3 | 2.7 | 3 | 2.83 | 2.91 | 3 | 3 | 2.8 | 2.9358 | 3.39 | 3.1 | 3 | 3.47 |  | 11
| 1 | 6 | 2.89 | 2.65 | 3.0701 | 2.6754 | 2.8885 | 3.9 | 3 | 2.7 | 2.98 | 2.83 | 2.91 | 3 | 3 | 3 | 2.982 | 3.28 | 3.03 | 3 | 2.9
| Mean |  | 3.0433 | 2.6483 | 3.047 | 2.9034 | 2.8885 | 3.9 | 3 | 2.6817 | 2.9883 | 2.84 | 2.8933 | 3 | 3 | 2.8833 | 2.982 | 3.3817 | 3.03 | 3 | 3.0867
| Median |  | 3 | 2.655 | 3.0507 | 2.8138 | 2.8902 | 3.9 | 3 | 2.69 | 2.99 | 2.83 | 2.9 | 3 | 3 | 2.9 | 2.9933 | 3.41 | 3.025 | 3 | 3.13
| Std Dev. |  | 0.16403 | 0.029944 | 0.027494 | 0.25329 | 0.042981 | 0.14142 | 0 | 0.02137 | 0.098472 | 0.030984 | 0.02582 | 0 | 0 | 0.075277 | 0.035238 | 0.097451 | 0.088091 | 0 | 0.45728
| Rel.Std.Dev. |  | 5.39% | 1.131% | 0.9023% | 8.724% | 1.488% | 3.626% | 0% | 0.7969% | 3.295% | 1.091% | 0.8924% | 0% | 0% | 2.611% | 1.182% | 2.882% | 2.907% | 0% | 14.81%
| PDM3 |  | 4.782% | -8.818% | 4.908% | -0.03457% | -0.5508% | 34.28% | 3.29% | -7.671% | 2.888% | -2.219% | -0.3828% | 3.29% | 3.29% | -0.7271% | 2.67% | 16.43% | 4.323% | 3.29% | 6.274%
| Z-Score (Absolute) |  | 0.44 | 2.41 | 0.47 | 0.57 | 0.67 | 6.62 | NA | 2.17 | 0.05 | 1.02 | 0.64 | NA | NA | 0.71 | 0 | 2.88 | 0.35 | NA | 0.75
Results from laboratories 9, 15, 16 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 466 (Certified Value 22.1 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 21.97 | 23.22 | 25.1877 |  | 20.4063 | 22.2 | 23.12 | 20.9 | 22.5 | 21.1 | 21 | 23.19 | 20.564 | 19.8 | 22.37 | 22.67 | 24 | 22.7 | 19.2 |  | 1
| 1 | 2 | 22.23 | 24.14 | 25.889 | 4.28 | 20.1335 | 21.9 | 22.66 | 21.4 | 23.2 | 20.9 | 19.35 | 22.33 | 21.2 | 22.3 | 22.4993 | 24.15 | 22.67 | 22.3 | 21 |  | 0
| 1 | 3 | 23.1 | 22.88 | 26.5335 | 4.43 | 19.9193 | 21.9 | 24.6 | 20.3 | 23.5 | 20.5 | 20.9 | 19.8 | 21.37 | 20.6 | 21.0661 | 22.73 | 23.93 | 20.7 | 20.1 |  | 16
| 1 | 4 | 22.64 | 21.9 | 25.0881 | 4.39 | 20.5757 | 21.8 | 25.13 | 20.6 | 22.8 | 20.8 | 20.1 | 22.15 | 21.094 | 21.4 | 21.1968 | 23.52 | 24.78 | 21 | 20.2 |  | 22.0859
| 1 | 5 | 23.11 | 22.92 | 26.075 | 3.54 | 19.5898 | 22.8 | 23.2 | 22.1 | 23.5 | 20.3 | 21.3 | 22.95 | 20.988 | 20.5 | 20.4384 | 22.31 | 23.74 | 20.2 | 20.3 |  | 12
| 1 | 6 | 22.97 | 23.75 | 26.8042 | 4.6 | 20.866 | 22.1 | 23.52 | 20.2 | 24.1 | 20.3 | 20.1 | 19.97 | 21.74 | 21.8 | 20.9072 | 23.87 | 24.41 | 23.2 | 20.8
| Mean |  | 22.67 | 23.135 | 25.9296 | 4.248 | 20.2484 | 22.1167 | 23.705 | 20.9167 | 23.2667 | 20.65 | 20.4583 | 21.7317 | 21.1593 | 21.0667 | 21.413 | 23.2083 | 23.9217 | 21.6833 | 20.2667
| Median |  | 22.805 | 23.07 | 25.982 | 4.39 | 20.2699 | 22 | 23.36 | 20.75 | 23.35 | 20.65 | 20.5 | 22.24 | 21.147 | 21 | 21.1314 | 23.125 | 23.965 | 21.65 | 20.25
| Std Dev. |  | 0.48021 | 0.77871 | 0.69442 | 0.41215 | 0.46235 | 0.3656 | 0.95448 | 0.72503 | 0.56804 | 0.33317 | 0.7324 | 1.4818 | 0.39275 | 0.9288 | 0.83291 | 0.74125 | 0.71848 | 1.2123 | 0.6314
| Rel.Std.Dev. |  | 2.118% | 3.366% | 2.678% | 9.702% | 2.283% | 1.653% | 4.026% | 3.466% | 2.441% | 1.613% | 3.58% | 6.819% | 1.856% | 4.409% | 3.89% | 3.194% | 3.003% | 5.591% | 3.115%
| PDM3 |  | 2.644% | 4.75% | 17.4% | -80.77% | -8.32% | 0.1391% | 7.331% | -5.294% | 5.346% | -6.502% | -7.369% | -1.604% | -4.195% | -4.615% | -3.047% | 5.082% | 8.312% | -1.823% | -8.237%
| Z-Score (Absolute) |  | 0.54 | 0.8 | 2.34 | 9.6 | 0.79 | 0.24 | 1.11 | 0.42 | 0.87 | 0.57 | 0.67 | 0.03 | 0.29 | 0.34 | 0.15 | 0.84 | 1.23 | 0 | 0.78
| Analytical results for Ca in OREAS 466 (Certified Value 0.226 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.2247 | 0.2675 | 0.22507 | 0.27051 | 0.2297 | 0.24 | 0.22 | 0.22 | 0.24 | 0.23 | 0.23 | 0.28 | 0.2264 | 0.22 | 0.324 | 0.2034 | 0.23 | 0.22 |  |  | 1
| 1 | 2 | 0.2261 | 0.2715 | 0.22565 | 0.26987 | 0.22502 | 0.22 | 0.22 | 0.22 | 0.23 | 0.23 | 0.22 | 0.26 | 0.2272 | 0.225 | 0.357 | 0.2181 | 0.24 | 0.24 |  |  | 0
| 1 | 3 | 0.2258 | 0.2718 | 0.22267 | 0.28417 | 0.22815 | 0.21 | 0.21 | 0.22 | 0.23 | 0.23 | 0.23 | 0.26 | 0.2288 | 0.225 | 0.34 | 0.2267 | 0.23 | 0.22 |  |  | 16
| 1 | 4 | 0.2245 | 0.264 | 0.22872 | 0.28264 | 0.2254 | 0.24 | 0.21 | 0.22 | 0.25 | 0.23 | 0.22 | 0.26 | 0.2212 | 0.224 | 0.339 | 0.207 | 0.23 | 0.24 |  |  | 0.22587
| 1 | 5 | 0.2253 | 0.2745 | 0.22659 | 0.2648 | 0.22647 | 0.23 | 0.21 | 0.22 | 0.24 | 0.23 | 0.23 | 0.26 | 0.2292 | 0.223 | 0.354 | 0.228 | 0.23 | 0.23 |  |  | 13
| 1 | 6 | 0.2243 | 0.2882 | 0.22699 | 0.25958 | 0.23146 | 0.23 | 0.22 | 0.22 | 0.24 | 0.23 | 0.23 | 0.26 | 0.2292 | 0.22 | 0.359 | 0.2086 | 0.23 | 0.22
| Mean |  | 0.22512 | 0.27292 | 0.22595 | 0.27193 | 0.2277 | 0.22833 | 0.215 | 0.22 | 0.23833 | 0.23 | 0.22667 | 0.26333 | 0.227 | 0.22283 | 0.3455 | 0.2153 | 0.23167 | 0.22833
| Median |  | 0.225 | 0.27165 | 0.22612 | 0.27019 | 0.22731 | 0.23 | 0.215 | 0.22 | 0.24 | 0.23 | 0.23 | 0.26 | 0.228 | 0.2235 | 0.347 | 0.21335 | 0.23 | 0.225
| Std Dev. |  | 0.00073326 | 0.0083449 | 0.0020388 | 0.0097387 | 0.0025413 | 0.01169 | 0.0054772 | 0 | 0.0075277 | 0 | 0.005164 | 0.008165 | 0.0030646 | 0.0023166 | 0.013576 | 0.010531 | 0.0040825 | 0.0098319
| Rel.Std.Dev. |  | 0.3257% | 3.058% | 0.9023% | 3.581% | 1.116% | 5.12% | 2.548% | 0% | 3.158% | 0% | 2.278% | 3.101% | 1.35% | 1.04% | 3.929% | 4.891% | 1.762% | 4.306%
| PDM3 |  | -0.3351% | 20.83% | 0.03282% | 20.39% | 0.8083% | 1.089% | -4.814% | -2.6% | 5.516% | 1.827% | 0.3511% | 16.58% | 0.4987% | -1.346% | 52.96% | -4.681% | 2.565% | 1.089%
| Z-Score (Absolute) |  | 0.44 | 6.86 | 0.32 | 6.7 | 0.05 | 0.05 | 1.99 | 1.22 | 1.58 | 0.3 | 0.21 | 5.39 | 0.16 | 0.79 | 17.94 | 1.94 | 0.56 | 0.05
| Analytical results for Cd in OREAS 466 (Indicative Value 0.04 ppm) | [21 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.04 | 0.09 | 0.1558 | < 0.2 | 0.7 | 0.03 | < 0.02 | < 0.02 | 0.02 | < 0.02 | 0.07 | < 0.1 | < 0.05 | 0.04797 | 0.04 | 0.03 | 3.16 |  |  |  | 1
| 1 | 2 | 0.05 | 0.06 | 0.1549 | < 0.2 | 0.7 | 0.03 | 0.02 | 0.02 | < 0.02 | < 0.02 | 0.07 | < 0.1 | < 0.05 | 0.05142 | 0.03 | 0.03 | 3.3 |  |  |  | 1
| 1 | 3 | 0.03 | 0.08 | 0.1863 | < 0.2 | 0.5 | < 0.02 | < 0.02 | < 0.02 | 0.03 | < 0.02 | 0.1 | < 0.1 | < 0.05 | 0.05586 | 0.04 | 0.08 | 3.14 |  |  |  | 16
| 1 | 4 | 0.05 | 0.09 | 0.1455 | < 0.2 | 0.4 | 0.03 | < 0.02 | < 0.02 | 0.03 | < 0.02 | 0.08 | < 0.1 | < 0.05 | 0.05219 | 0.05 | 0.04 | 3.36 |  |  |  | 0.040178
| 1 | 5 | 0.03 | 0.05 | 0.1201 | < 0.2 | 0.4 | < 0.02 | 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.07 | < 0.1 | < 0.05 | 0.05603 | 0.04 | 0.02 | 3.2 |  |  |  | 7
| 1 | 6 | 0.03 | 0.05 | 0.1182 | < 0.2 | 0.6 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.02 | 0.1 | < 0.1 | < 0.05 | 0.05829 | 0.04 | 0.05 | 3.16
| Mean |  | 0.038333 | 0.07 | 0.1468 |  | 0.55 | 0.03 | 0.02 | 0.02 | 0.026667 | 0.02 | 0.081667 |  |  | 0.053627 | 0.04 | 0.041667 | 3.22
| Median |  | 0.035 | 0.07 | 0.1502 |  | 0.55 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 | 0.075 |  |  | 0.054025 | 0.04 | 0.035 | 3.18
| Std Dev. |  | 0.0098319 | 0.018974 | 0.025453 |  | 0.13784 | 0 | 0 |  | 0.0057735 |  | 0.01472 |  |  | 0.00378 | 0.0063246 | 0.02137 | 0.089443
| Rel.Std.Dev. |  | 25.65% | 27.11% | 17.34% |  | 25.06% | 0% | 0% |  | 21.65% |  | 18.02% |  |  | 7.049% | 15.81% | 51.29% | 2.778%
| PDM3 |  | -4.592% | 74.22% | 265.4% |  | 1269% | -25.33% | -50.22% | -50.22% | -33.63% | -50.22% | 103.3% |  |  | 33.47% | -0.4435% | 3.705% | 7914%
| Z-Score (Absolute) |  | 0.08 | 0.67 | 2.5 | 1.39 | 12.1 | 0.52 | 0.67 | 0.71 | 0.56 | 0.71 | 0.95 | 0.2 | 0.4 | 0.28 | 0.04 | 0 | 75.64
| Analytical results for Ce in OREAS 466 (Certified Value 339 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 0
0
| 1 | 1 | 345.7 | 333.5 | 402.948 | 368.766 | 324.5417 | 356 | 315.04 | 344.74 | 330 | 361 | 326 | 345 | 411.54 | 357 | 328 | 311.8733 | 411 | 331.75 | 413 | 193 | 1
| 1 | 2 | 332.6 | 335.5 | 402.17 | 361.479 | 336.7105 | 339 | 311.15 | 342.08 | 335 | 350 | 320 | 339 | 387.59 | 351 | 335 | 294.7448 | 438 | 328.24 | 411 | 205 | 0
| 1 | 3 | 345.5 | 332.9 | 402.0152 | 366.603 | 329.0702 | 323 | 311.96 | 338.35 | 324 | 352 | 313 | 334 | 376.07 | 356 | 336 | 298.0431 | 407 | 332.73 | 404 | 184 | 16
| 1 | 4 | 338.9 | 322.1 | 401.3554 | 362.4935 | 336.0978 | 338 | 310.41 | 337.64 | 324 | 373 | 323 | 332 | 392.59 | 351 | 341 | 300.7101 | 414 | 333.91 | 406 | 191 | 339.3862
| 1 | 5 | 352.8 | 332.6 | 403.138 | 362.993 | 336.5389 | 352 | 315.65 | 330.97 | 335 | 361 | 326 | 344 | 391.71 | 354 | 327 | 319.2416 | 405 | 328.3 | 403 | 200 | 14
| 1 | 6 | 343.9 | 335.6 | 402.0766 | 361.425 | 337.5776 | 336 | 312.14 | 348.77 | 329 | 378 | 321 | 330 | 377.25 | 349 | 341 | 319.6307 | 416 | 341.85 | 426 | 196
| Mean |  | 343.2333 | 332.0333 | 402.2839 | 363.9599 | 333.4228 | 340.6667 | 312.725 | 340.425 | 329.5 | 362.5 | 321.5 | 337.3333 | 389.4583 | 353 | 334.6667 | 307.3739 | 415.1667 | 332.7967 | 410.5 | 194.8333
| Median |  | 344.7 | 333.2 | 402.1233 | 362.7432 | 336.3183 | 338.5 | 312.05 | 340.215 | 329.5 | 361 | 322 | 336.5 | 389.65 | 352.5 | 335.5 | 306.2917 | 412.5 | 332.24 | 408.5 | 194.5
| Std Dev. |  | 6.8561 | 5.0326 | 0.6574 | 3.0249 | 5.3433 | 11.894 | 2.1299 | 6.2041 | 4.9295 | 11.149 | 4.8477 | 6.314 | 12.9171 | 3.1623 | 6.0882 | 10.9752 | 11.9234 | 5.0057 | 8.5499 | 7.3052
| Rel.Std.Dev. |  | 1.998% | 1.516% | 0.1634% | 0.8311% | 1.603% | 3.491% | 0.6811% | 1.822% | 1.496% | 3.076% | 1.508% | 1.872% | 3.317% | 0.8958% | 1.819% | 3.571% | 2.872% | 1.504% | 2.083% | 3.749%
| PDM3 |  | 1.134% | -2.167% | 18.53% | 7.241% | -1.757% | 0.3773% | -7.856% | 0.3061% | -2.913% | 6.81% | -5.27% | -0.6049% | 14.75% | 4.011% | -1.391% | -9.432% | 22.33% | -1.942% | 20.95% | -42.59%
| Z-Score (Absolute) |  | 0.19 | 0.29 | 2.71 | 1.07 | 0.23 | 0.08 | 1.12 | 0.07 | 0.4 | 1.01 | 0.74 | 0.07 | 2.17 | 0.6 | 0.18 | 1.35 | 3.27 | 0.26 | 3.07 | 6.17
| Analytical results for Co in OREAS 466 (Certified Value 2.92 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.8 | 3.7 | 2.508 | 4.32 | 2.8595 | 3 | 2.7 | 3 | 2.8 | 3 | 2.9 | 3 | 3.1 | 3 | 2.895 | 3.2 | 3 | 2.9 |  |  | 1
| 1 | 2 | 3 | 3.4 | 2.5782 | 1.76 | 2.771 | 3 | 2.7 | 3 | 2.7 | 2.9 | 2.9 | 2.9 | 3.2 | 3 | 3.0002 | 3 | 3 | 2.9 |  |  | 0
| 1 | 3 | 3 | 3.4 | 2.5082 | 5.95 | 2.8596 | 3 | 2.8 | 3.1 | 2.9 | 2.9 | 2.9 | 2.7 | 3.1 | 3 | 2.9701 | 3 | 3 | 2.9 |  |  | 16
| 1 | 4 | 3 | 3.5 | 2.4981 | 4 | 2.8321 | 3 | 2.8 | 3 | 2.8 | 2.9 | 2.9 | 2.8 | 3.2 | 3 | 3.0897 | 3.1 | 2.8 | 2.9 |  |  | 2.9229
| 1 | 5 | 2.9 | 3.5 | 2.5594 | 25.7 | 2.8939 | 3 | 2.7 | 3.1 | 2.9 | 2.9 | 2.9 | 3.1 | 3 | 3 | 2.842 | 2.9 | 2.8 | 2.9 |  |  | 15
| 1 | 6 | 3 | 3.4 | 2.5338 |  | 2.8406 | 2 | 2.8 | 2.9 | 2.8 | 2.9 | 2.8 | 2.7 | 3.1 | 3 | 2.9342 | 3 | 3.1 | 2.9
| Mean |  | 2.95 | 3.4833 | 2.531 | 8.346 | 2.8428 | 2.8333 | 2.75 | 3.0167 | 2.8167 | 2.9167 | 2.8833 | 2.8667 | 3.1167 | 3 | 2.9552 | 3.0333 | 2.95 | 2.9
| Median |  | 3 | 3.45 | 2.521 | 4.32 | 2.8501 | 3 | 2.75 | 3 | 2.8 | 2.9 | 2.9 | 2.85 | 3.1 | 3 | 2.9521 | 3 | 3 | 2.9
| Std Dev. |  | 0.083666 | 0.1169 | 0.032188 | 9.8155 | 0.041107 | 0.40825 | 0.054772 | 0.075277 | 0.075277 | 0.040825 | 0.040825 | 0.1633 | 0.075277 | 0 | 0.08628 | 0.10328 | 0.12247 | 0
| Rel.Std.Dev. |  | 2.836% | 3.356% | 1.272% | 117.6% | 1.446% | 14.41% | 1.992% | 2.495% | 2.673% | 1.4% | 1.416% | 5.696% | 2.415% | 0% | 2.92% | 3.405% | 4.152% | 0%
| PDM3 |  | 0.9265% | 19.17% | -13.41% | 185.5% | -2.742% | -3.065% | -5.916% | 3.207% | -3.635% | -0.2139% | -1.354% | -1.925% | 6.629% | 2.637% | 1.104% | 3.778% | 0.9265% | -0.7841%
| Z-Score (Absolute) |  | 0.13 | 4.27 | 3.12 | 41.98 | 0.7 | NA | 1.42 | 0.65 | 0.9 | 0.13 | 0.39 | 0.52 | 1.42 | NA | 0.17 | 0.78 | 0.13 | 0.26
Results from laboratories 7 and 17 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 466 (Certified Value 132 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 150 | 139.2 | 156.9458 | 101.2605 | 129.4981 | 117 | 126 | 131 | 133 | 130 | 132 | 128 | 114 | 130 | 151.1851 | 168 | 137 | 122 | 126 |  | 1
| 1 | 2 | 152 | 141.1 | 157.4294 | 106.4295 | 131.8629 | 112 | 125 | 130 | 133 | 131 | 129 | 126 | 112 | 129 | 156.6974 | 176 | 140 | 123 | 127 |  | 0
| 1 | 3 | 149 | 141.3 | 154.24 | 112.0705 | 133.461 | 103 | 129 | 129 | 135 | 124 | 133 | 125 | 114 | 136 | 157.1513 | 187 | 137 | 122 | 123 |  | 16
| 1 | 4 | 146 | 137.2 | 157.3466 | 112.5184 | 133.3149 | 133 | 122 | 131 | 141 | 127 | 129 | 125 | 110 | 130 | 162.8159 | 164 | 166 | 123 | 127 |  | 131.7938
| 1 | 5 | 150 | 141 | 155.328 | 102.4832 | 131.9343 | 97 | 127 | 129 | 133 | 130 | 132 | 129 | 118 | 133 | 153.2833 | 175 | 156 | 123 | 127 |  | 16
| 1 | 6 | 150 | 139.7 | 152.6912 | 110.0005 | 137.0812 | 104 | 127 | 130 | 134 | 130 | 132 | 118 | 113 | 138 | 159.2003 | 194 | 154 | 121 | 129
| Mean |  | 149.5 | 139.9167 | 155.6635 | 107.4604 | 132.8587 | 111 | 126 | 130 | 134.8333 | 128.6667 | 131.1667 | 125.1667 | 113.5 | 132.6667 | 156.7222 | 177.3333 | 148.3333 | 122.3333 | 126.5
| Median |  | 150 | 140.35 | 156.1369 | 108.215 | 132.6246 | 108 | 126.5 | 130 | 133.5 | 130 | 132 | 125.5 | 113.5 | 131.5 | 156.9244 | 175.5 | 147 | 122.5 | 127
| Std Dev. |  | 1.9748 | 1.5766 | 1.9271 | 4.849 | 2.5116 | 12.8841 | 2.3664 | 0.89443 | 3.1252 | 2.6583 | 1.7224 | 3.8687 | 2.6646 | 3.6697 | 4.1459 | 11.3431 | 12.0775 | 0.8165 | 1.9748
| Rel.Std.Dev. |  | 1.321% | 1.127% | 1.238% | 4.512% | 1.89% | 11.61% | 1.878% | 0.688% | 2.318% | 2.066% | 1.313% | 3.091% | 2.348% | 2.766% | 2.645% | 6.397% | 8.142% | 0.6674% | 1.561%
| PDM3 |  | 13.43% | 6.163% | 18.11% | -18.46% | 0.808% | -15.78% | -4.396% | -1.361% | 2.306% | -2.373% | -0.4758% | -5.028% | -13.88% | 0.6623% | 18.91% | 34.55% | 12.55% | -7.178% | -4.017%
| Z-Score (Absolute) |  | 1.41 | 0.67 | 1.89 | 1.83 | 0.13 | 1.55 | 0.4 | 0.09 | 0.28 | 0.19 | 0 | 0.46 | 1.36 | 0.12 | 1.97 | 3.56 | 1.32 | 0.68 | 0.36
| Analytical results for Cs in OREAS 466 (Certified Value 10.8 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 10.97 | 11.7 | 9.908 | 11.1203 | 10.6 | 10.5 | 9.96 | 11.95 | 10.4 | 10.75 | 11.4 | 10.88 | 10.6 | 10.2532 | 10.92 | 10.78 | 13 | 10.7 |  |  | 1
| 1 | 2 | 11.04 | 12.1 | 9.602 | 11.2269 | 10.2 | 10.4 | 10.05 | 11.15 | 10.2 | 10.55 | 11 | 10.68 | 10.8 | 9.0627 | 11.73 | 10.73 | 13 | 10.8 |  |  | 0
| 1 | 3 | 11.11 | 11.8 | 9.436 | 11.3468 | 10.1 | 10.5 | 9.99 | 11.45 | 10.15 | 10.45 | 10 | 10.8 | 10.6 | 8.8242 | 11.05 | 10.78 | 12 | 10.7 |  |  | 16
| 1 | 4 | 11.07 | 10.6 | 10 | 11.2578 | 10.3 | 10.3 | 9.74 | 12 | 10.4 | 10.5 | 10.7 | 10.7 | 10.8 | 8.6145 | 11.03 | 10.88 | 12 | 11 |  |  | 10.7543
| 1 | 5 | 10.82 | 11.9 | 9.288 | 11.1426 | 10.6 | 10.2 | 10.35 | 11.5 | 10.6 | 11 | 11.6 | 10.7 | 10.7 | 9.5286 | 10.92 | 10.82 | 12 | 10.9 |  |  | 17
| 1 | 6 | 11.23 | 11.4 | 9.812 | 11.6057 | 10.4 | 10.8 | 9.95 | 11.9 | 10.35 | 10.45 | 10.3 | 10.96 | 10.8 | 10.3771 | 11 | 10.98 | 13 | 10.5
| Mean |  | 11.04 | 11.5833 | 9.6743 | 11.2833 | 10.3667 | 10.45 | 10.0067 | 11.6583 | 10.35 | 10.6167 | 10.8333 | 10.7867 | 10.7167 | 9.4434 | 11.1083 | 10.8283 | 12.5 | 10.7667
| Median |  | 11.055 | 11.75 | 9.707 | 11.2423 | 10.35 | 10.45 | 9.975 | 11.7 | 10.375 | 10.525 | 10.85 | 10.75 | 10.75 | 9.2957 | 11.015 | 10.8 | 12.5 | 10.75
| Std Dev. |  | 0.13799 | 0.53448 | 0.27961 | 0.17785 | 0.20656 | 0.20736 | 0.19826 | 0.34266 | 0.16125 | 0.21833 | 0.62183 | 0.11431 | 0.098319 | 0.74165 | 0.30941 | 0.089536 | 0.54772 | 0.17512
| Rel.Std.Dev. |  | 1.25% | 4.614% | 2.89% | 1.576% | 1.993% | 1.984% | 1.981% | 2.939% | 1.558% | 2.056% | 5.74% | 1.06% | 0.9174% | 7.854% | 2.785% | 0.8269% | 4.382% | 1.626%
| PDM3 |  | 2.656% | 7.709% | -10.04% | 4.919% | -3.605% | -2.83% | -6.952% | 8.406% | -3.76% | -1.28% | 0.7346% | 0.3007% | -0.3502% | -12.19% | 3.292% | 0.6881% | 16.23% | 0.1147%
| Z-Score (Absolute) |  | 0.54 | 1.61 | 2.16 | 1.02 | 0.79 | 0.62 | 1.5 | 1.76 | 0.82 | 0.3 | 0.13 | 0.04 | 0.1 | 2.61 | 0.67 | 0.12 | NA | 0
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 466 (Certified Value 34.1 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 34.2 | 24.9 | 40.024 | 55.61 | 34.1439 | 36 | 34.9 | 32.2 | 32.9 | 34.7 | 33.7 | 35.5 | 31.2 | 34 | 33.7644 | 39.8 | 34.7 | 33.2 | 38 |  | 1
| 1 | 2 | 33.9 | 25.3 | 40.4938 | 54.19 | 33.283 | 35 | 34.5 | 32.3 | 30.5 | 33.4 | 34 | 36 | 31.6 | 35 | 34.6327 | 49.4 | 33.9 | 33.6 | 35 |  | 0
| 1 | 3 | 34.2 | 22.2 | 40.3063 | 55.02 | 34.3565 | 37 | 34.6 | 32.1 | 32.7 | 33.2 | 33.2 | 33.8 | 31.7 | 35 | 34.1432 | 51.6 | 33.7 | 33.2 | 36 |  | 16
| 1 | 4 | 34.3 | 21.7 | 40.2839 | 62.22 | 34.6102 | 36 | 35.2 | 32.3 | 35.8 | 34 | 34 | 34 | 31.365 | 35 | 36.241 | 45 | 34.4 | 33.8 | 37 |  | 34.1434
| 1 | 5 | 34.5 | 22.5 | 40.2157 | 56.78 | 34.8395 | 35 | 34.5 | 32.8 | 35 | 33.3 | 34.1 | 36.3 | 31.5 | 34 | 33.686 | 39.8 | 34.3 | 33.5 | 37 |  | 18
| 1 | 6 | 34.9 | 24.8 | 40.673 | 55.42 | 34.6056 | 35 | 35 | 31.8 | 34.3 | 33.8 | 33 | 33.9 | 31.7 | 35 | 34.1375 | 51.6 | 34.6 | 33.6 | 39
| Mean |  | 34.3333 | 23.5667 | 40.3328 | 56.54 | 34.3065 | 35.6667 | 34.7833 | 32.25 | 33.5333 | 33.7333 | 33.6667 | 34.9167 | 31.5108 | 34.6667 | 34.4341 | 46.2 | 34.2667 | 33.4833 | 37
| Median |  | 34.25 | 23.65 | 40.2951 | 55.515 | 34.4811 | 35.5 | 34.75 | 32.25 | 33.6 | 33.6 | 33.85 | 34.75 | 31.55 | 35 | 34.1404 | 47.2 | 34.35 | 33.55 | 37
| Std Dev. |  | 0.33862 | 1.5996 | 0.22532 | 2.9077 | 0.5554 | 0.8165 | 0.29269 | 0.32711 | 1.9065 | 0.56451 | 0.46332 | 1.1444 | 0.19876 | 0.5164 | 0.94712 | 5.5122 | 0.39328 | 0.24014 | 1.4142
| Rel.Std.Dev. |  | 0.9863% | 6.787% | 0.5586% | 5.143% | 1.619% | 2.289% | 0.8415% | 1.014% | 5.685% | 1.673% | 1.376% | 3.278% | 0.6308% | 1.49% | 2.751% | 11.93% | 1.148% | 0.7172% | 3.822%
| PDM3 |  | 0.5562% | -30.98% | 18.13% | 65.6% | 0.4775% | 4.461% | 1.874% | -5.546% | -1.787% | -1.201% | -1.396% | 2.265% | -7.71% | 1.532% | 0.8514% | 35.31% | 0.3609% | -1.933% | 8.366%
| Z-Score (Absolute) |  | 0 | 9.08 | 5.06 | 18.72 | 0.02 | 1.12 | 0.38 | 1.76 | 0.67 | 0.51 | 0.56 | 0.49 | 2.38 | 0.28 | 0.08 | 10 | 0.06 | 0.72 | 2.25
| Analytical results for Dy in OREAS 466 (Certified Value 5.57 ppm) | [21 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 5.8 | 6.8331 | 7.06 | 5.7053 | 6.6 | 5.39 | 5.6 | < 100 | 5.48 | 5.4 | 5.68 | 7.1081 | 5.37 | 5.78 | 5.87 | [5 empty] | 1
| 1 | 2 | 5.3 | 6.7521 | 6.958 | 5.6149 | 6.1 | 5.35 | 5.6 | < 100 | 5.47 | 5.5 | 5.68 | 7.062 | 5.56 | 5.79 | 5.62 | [5 empty] | 0
| 1 | 3 | 5 | 6.6707 | 6.993 | 5.675 | 6 | 5.34 | 5.5 | < 100 | 5.45 | 4.9 | 5.5 | 7.2608 | 5.05 | 5.98 | 5.63 | [5 empty] | 16
| 1 | 4 | 5.4 | 6.8061 | 6.546 | 5.4413 | 6.1 | 5.28 | 5.5 | < 100 | 5.55 | 5.3 | 5.56 | 7.3883 | 5.3 | 5.75 | 5.44 | [5 empty] | 5.5683
| 1 | 5 | 5.4 | 6.8069 | 6.585 | 5.8309 | 6.6 | 5.42 | 5.5 | < 100 | 5.72 | 5.5 | 5.64 | 7.1519 | 5.3 | 5.61 | 5.36 | [5 empty] | 19
| 1 | 6 | 5.4 | 6.7388 | 6.754 | 5.6317 | 6.2 | 5.37 | 5.7 | < 100 | 5.42 | 5.1 | 5.6 | 7.0217 | 5.35 | 5.95 | 6
| Mean |  | 5.3833 | 6.768 | 6.816 | 5.6498 | 6.2667 | 5.3583 | 5.5667 |  | 5.515 | 5.2833 | 5.61 | 7.1655 | 5.3217 | 5.81 | 5.6533
| Median |  | 5.4 | 6.7791 | 6.856 | 5.6533 | 6.15 | 5.36 | 5.55 |  | 5.475 | 5.35 | 5.62 | 7.13 | 5.325 | 5.785 | 5.625
| Std Dev. |  | 0.25626 | 0.059614 | 0.21961 | 0.12768 | 0.26583 | 0.047924 | 0.08165 |  | 0.10932 | 0.24014 | 0.071274 | 0.1368 | 0.16412 | 0.13667 | 0.24509
| Rel.Std.Dev. |  | 4.76% | 0.8808% | 3.222% | 2.26% | 4.242% | 0.8944% | 1.467% |  | 1.982% | 4.545% | 1.27% | 1.909% | 3.084% | 2.352% | 4.335%
| PDM3 |  | -3.322% | 21.54% | 22.41% | 1.464% | 12.54% | -3.771% | -0.02966% |  | -0.9575% | -5.118% | 0.7485% | 28.68% | -4.43% | 4.34% | 1.527%
| Z-Score (Absolute) |  | 0.62 | 2.59 | 2.7 | 0 | 1.43 | 0.67 | 0.19 | 102.59 | 0.31 | 0.85 | 0.09 | 3.51 | 0.76 | 0.37 | 0.01
| Analytical results for Er in OREAS 466 (Certified Value 2.19 ppm) | [21 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 2.19 | 2.8556 | 2.748 | 2.1658 | 2.9 | 1.89 | 2 | 2.13 | 2 | 1.93 | 3.4448 | 1.86 | 2.23 | 2.04 | [6 empty] | 1
| 1 | 2 | 2.2 | 2.9206 | 2.751 | 2.0788 | 2.8 | 1.96 | 2 | 2.18 | 2 | 1.9568 | 3.4903 | 1.98 | 2.13 | 1.99 | [6 empty] | 0
| 1 | 3 | 2.26 | 2.7874 | 2.586 | 2.1969 | 2.5 | 1.87 | 2 | 2.16 | 1.8 | 1.9459 | 3.4974 | 1.8 | 2.2 | 2.01 | [6 empty] | 16
| 1 | 4 | 2.23 | 2.854 | 2.25 | 2.1907 | 2.6 | 1.85 | 2 | 2.15 | 1.9 | 1.92 | 3.596 | 1.86 | 2.07 | 1.97 | [6 empty] | 2.1932
| 1 | 5 | 2.28 | 2.8308 | 2.371 | 2.1338 | 2.7 | 1.83 | 2 | 2.23 | 2.1 | 1.92 | 3.4697 | 1.86 | 2.11 | 1.92 | [6 empty] | 20
| 1 | 6 | 2.25 | 2.8325 | 2.863 | 2.1488 | 2.7 | 1.94 | 2.1 | 2.06 | 1.9 | 1.92 | 3.442 | 1.88 | 2.16 | 2.18
| Mean |  | 2.235 | 2.8468 | 2.5948 | 2.1525 | 2.7 | 1.89 | 2.0167 | 2.1517 | 1.95 | 1.9321 | 3.49 | 1.8733 | 2.15 | 2.0183
| Median |  | 2.24 | 2.8432 | 2.667 | 2.1573 | 2.7 | 1.88 | 2 | 2.155 | 1.95 | 1.925 | 3.48 | 1.86 | 2.145 | 2
| Std Dev. |  | 0.035071 | 0.043733 | 0.24035 | 0.04338 | 0.14142 | 0.05099 | 0.040825 | 0.056362 | 0.10488 | 0.015769 | 0.056659 | 0.058878 | 0.058992 | 0.088863
| Rel.Std.Dev. |  | 1.569% | 1.536% | 9.263% | 2.015% | 5.238% | 2.698% | 2.024% | 2.619% | 5.379% | 0.8162% | 1.623% | 3.143% | 2.744% | 4.403%
| PDM3 |  | 1.907% | 29.8% | 18.31% | -1.856% | 23.11% | -13.82% | -8.048% | -1.892% | -11.09% | -11.9% | 59.13% | -14.58% | -1.968% | -7.972%
| Z-Score (Absolute) |  | 0.27 | 2.24 | 1.43 | 0.01 | 1.77 | 0.84 | 0.43 | 0 | 0.65 | 0.7 | 4.3 | 0.89 | 0 | 0.43
| Analytical results for Eu in OREAS 466 (Certified Value 4.26 ppm) | [21 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 4.7 | 5.3793 | 5.365 | 4.3233 | 4.71 | 4.32 | 4.3 | 4.26 | 4.4 | 4.22 | 4.1311 | 4.23 | 4.44 | 3.88 | [6 empty] | 1
| 1 | 2 | 4.1 | 5.3155 | 4.873 | 4.3102 | 4.28 | 3.94 | 4.2 | 4.08 | 4.4 | 4.2 | 4.1122 | 4.42 | 4.42 | 3.86 | [6 empty] | 0
| 1 | 3 | 4.3 | 5.3747 | 4.956 | 4.493 | 4.21 | 4.07 | 4.2 | 4.15 | 3.9 | 4.2 | 4.1446 | 3.97 | 4.46 | 3.68 | [6 empty] | 16
| 1 | 4 | 4.4 | 5.358 | 4.858 | 4.3331 | 4.35 | 3.94 | 4.2 | 3.97 | 4.3 | 4.22 | 4.2052 | 4.17 | 4.35 | 3.57 | [6 empty] | 4.2557
| 1 | 5 | 4.5 | 5.4844 | 4.545 | 4.3022 | 4.5 | 4.29 | 4.2 | 4.25 | 4.4 | 4.28 | 4.165 | 4 | 4.39 | 3.51 | [6 empty] | 21
| 1 | 6 | 4.3 | 5.4296 | 5.013 | 4.428 | 4.4 | 4.13 | 4.3 | 3.98 | 4 | 4.2 | 4.1487 | 4.17 | 4.51 | 3.97
| Mean |  | 4.3833 | 5.3902 | 4.935 | 4.365 | 4.4083 | 4.115 | 4.2333 | 4.115 | 4.2333 | 4.22 | 4.1511 | 4.16 | 4.4283 | 3.745
| Median |  | 4.35 | 5.377 | 4.9145 | 4.3282 | 4.375 | 4.1 | 4.2 | 4.115 | 4.35 | 4.21 | 4.1466 | 4.17 | 4.43 | 3.77
| Std Dev. |  | 0.20412 | 0.058986 | 0.26582 | 0.077521 | 0.17815 | 0.16502 | 0.05164 | 0.12724 | 0.22509 | 0.030984 | 0.031909 | 0.16395 | 0.055648 | 0.18555
| Rel.Std.Dev. |  | 4.657% | 1.094% | 5.386% | 1.776% | 4.041% | 4.01% | 1.22% | 3.092% | 5.317% | 0.7342% | 0.7687% | 3.941% | 1.257% | 4.955%
| PDM3 |  | 2.999% | 26.66% | 15.96% | 2.568% | 3.586% | -3.306% | -0.5257% | -3.306% | -0.5257% | -0.839% | -2.457% | -2.249% | 4.056% | -12%
| Z-Score (Absolute) |  | 0.81 | 6.24 | 3.79 | 0.71 | 0.94 | 0.64 | 0 | 0.64 | 0 | 0.07 | 0.44 | 0.4 | 1.05 | 2.63
| Analytical results for Fe in OREAS 466 (Certified Value 4.97 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 5.05 | 4.7642 | 5.1316 | 6.4544 | 5.1645 | 5.18 | 4.93 | 4.64 | 5.08 | 4.91 | 4.88 | 5.13 | 5.08 | 5.01 | 4.43 | 5.16 | 5.03 | 5.13 |  |  | 1
| 1 | 2 | 5.06 | 4.855 | 5.1279 | 6.4002 | 4.9656 | 4.92 | 4.92 | 4.63 | 4.83 | 4.84 | 4.77 | 5.17 | 4.9 | 5.05 | 4.96 | 5.04 | 5.05 | 5.16 |  |  | 0
| 1 | 3 | 5.01 | 4.8207 | 5.1835 | 6.3399 | 5.1816 | 4.85 | 4.84 | 4.52 | 4.94 | 4.77 | 4.85 | 4.92 | 5.04 | 5.01 | 5.17 | 5.13 | 5.02 | 5.02 |  |  | 16
| 1 | 4 | 4.89 | 4.6856 | 5.12 | 6.4804 | 5.1456 | 5.09 | 4.88 | 4.56 | 5.32 | 4.9 | 4.78 | 4.89 | 4.9 | 5.05 | 4.74 | 5.18 | 5.01 | 5.27 |  |  | 4.9673
| 1 | 5 | 4.94 | 4.772 | 5.1211 | 6.3144 | 4.9681 | 5.12 | 4.9 | 4.64 | 5.09 | 4.84 | 4.85 | 5.02 | 4.96 | 4.94 | 4.35 | 5.11 | 5.08 | 5.08 |  |  | 22
| 1 | 6 | 5.02 | 4.7663 | 5.1341 | 6.2523 | 5.2007 | 5.01 | 4.98 | 4.63 | 5.09 | 4.85 | 4.83 | 4.95 | 4.92 | 4.99 | 4.93 | 5.17 | 5.05 | 5.23
| Mean |  | 4.995 | 4.7773 | 5.1364 | 6.3736 | 5.1043 | 5.0283 | 4.9083 | 4.6033 | 5.0583 | 4.8517 | 4.8267 | 5.0133 | 4.9667 | 5.0083 | 4.7633 | 5.1317 | 5.04 | 5.1483
| Median |  | 5.015 | 4.7691 | 5.1298 | 6.3701 | 5.155 | 5.05 | 4.91 | 4.63 | 5.085 | 4.845 | 4.84 | 4.985 | 4.94 | 5.01 | 4.835 | 5.145 | 5.04 | 5.145
| Std Dev. |  | 0.066558 | 0.057707 | 0.023741 | 0.087187 | 0.10808 | 0.12576 | 0.047504 | 0.050859 | 0.16558 | 0.050365 | 0.043205 | 0.11501 | 0.076594 | 0.041191 | 0.32073 | 0.051929 | 0.025298 | 0.092826
| Rel.Std.Dev. |  | 1.332% | 1.208% | 0.4622% | 1.368% | 2.117% | 2.501% | 0.9678% | 1.105% | 3.273% | 1.038% | 0.8951% | 2.294% | 1.542% | 0.8224% | 6.733% | 1.012% | 0.5019% | 1.803%
| PDM3 |  | 0.5578% | -3.825% | 3.404% | 28.31% | 2.759% | 1.229% | -1.187% | -7.327% | 1.833% | -2.328% | -2.831% | 0.9269% | -0.01262% | 0.8262% | -4.106% | 3.309% | 1.464% | 3.645%
| Z-Score (Absolute) |  | 0.1 | 1.41 | 0.76 | 8.23 | 0.56 | 0.11 | 0.62 | 2.46 | 0.29 | 0.96 | 1.11 | 0.02 | 0.27 | 0.02 | 1.49 | 0.73 | 0.18 | 0.83
| Analytical results for Ga in OREAS 466 (Certified Value 22.4 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 22.3 | 18.75 | 24.9856 | 22.71 | 21.4623 | 20.9 | 20.78 | 21.7 | 23.4 | 22.2 | 22.9 | 25.71 | 21.6 | 22.441 | 22.25 | 15.8 | 30 |  |  |  | 1
| 1 | 2 | 22.5 | 19.84 | 24.7882 | 22.34 | 21.318 | 20 | 20.84 | 22 | 20.7 | 21.7 | 23 | 25.58 | 21.9 | 22.1624 | 22.36 | 16 | 30 |  |  |  | 0
| 1 | 3 | 20.8 | 18.74 | 24.2791 | 23.08 | 21.9151 | 19.6 | 20.84 | 22 | 22.7 | 21.5 | 22.5 | 24.42 | 21.8 | 22.475 | 22.67 | 15.1 | 30 |  |  |  | 16
| 1 | 4 | 21.1 | 17.04 | 24.9865 | 22.77 | 21.6611 | 20.4 | 21 | 21.7 | 22.8 | 22 | 22.7 | 23.84 | 22.2 | 23.17 | 22.56 | 15 | 30 |  |  |  | 22.3564
| 1 | 5 | 22.2 | 19.01 | 24.7746 | 22.05 | 21.787 | 21.4 | 20.77 | 22.5 | 24 | 22.1 | 22.7 | 25.46 | 21.4 | 22.1831 | 22.46 | 15.1 | 30 |  |  |  | 23
| 1 | 6 | 21.7 | 18 | 24.5268 | 22.74 | 22.1704 | 20.2 | 20.88 | 21.6 | 24.5 | 21.9 | 22.7 | 23.71 | 21.6 | 22.3878 | 22.38 | 15.9 | 30
| Mean |  | 21.7667 | 18.5633 | 24.7234 | 22.615 | 21.719 | 20.4167 | 20.8517 | 21.9167 | 23.0167 | 21.9 | 22.75 | 24.7867 | 21.75 | 22.4699 | 22.4467 | 15.4833 | 30
| Median |  | 21.95 | 18.745 | 24.7813 | 22.725 | 21.724 | 20.3 | 20.84 | 21.85 | 23.1 | 21.95 | 22.7 | 24.94 | 21.7 | 22.4144 | 22.42 | 15.45 | 30
| Std Dev. |  | 0.69186 | 0.95219 | 0.27599 | 0.36325 | 0.30883 | 0.64627 | 0.083526 | 0.33116 | 1.3288 | 0.26077 | 0.17607 | 0.90831 | 0.28107 | 0.3673 | 0.15069 | 0.46224 | 0
| Rel.Std.Dev. |  | 3.179% | 5.129% | 1.116% | 1.606% | 1.422% | 3.165% | 0.4006% | 1.511% | 5.773% | 1.191% | 0.7739% | 3.665% | 1.292% | 1.635% | 0.6713% | 2.985% | 0%
| PDM3 |  | -2.638% | -16.97% | 10.59% | 1.157% | -2.851% | -8.676% | -6.73% | -1.967% | 2.954% | -2.041% | 1.761% | 10.87% | -2.712% | 0.5078% | 0.404% | -30.74% | 34.19%
| Z-Score (Absolute) |  | 0.12 | 2.71 | 2.27 | 0.57 | 0.16 | 1.21 | 0.86 | 0 | 0.89 | 0.01 | 0.67 | 2.32 | 0.13 | 0.45 | 0.43 | 5.21 | 6.54
| Analytical results for Gd in OREAS 466 (Certified Value 11.4 ppm) | [21 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 11.6 | 18.3627 | 14.79 | 11.2261 | 13 | 10.96 | 11.2 | < 100 | 9.64 | 13.1 | 11.92 | 11.3688 | 12.59 | 11.14 | 11 | [5 empty] | 1
| 1 | 2 | 11.2 | 18.145 | 14.31 | 11.0444 | 12.2 | 10.49 | 11.2 | < 100 | 9.85 | 12.8 | 11.84 | 11.0503 | 13.13 | 11.05 | 10.7 | [5 empty] | 0
| 1 | 3 | 11.5 | 18.4672 | 14.39 | 11.4852 | 12.1 | 10.78 | 11.1 | < 100 | 9.73 | 11.6 | 11.6 | 11.3786 | 11.67 | 12.05 | 10.3 | [5 empty] | 16
| 1 | 4 | 11.3 | 18.2902 | 15.69 | 11.2615 | 12.5 | 10.8 | 11 | < 100 | 9.71 | 12.7 | 11.68 | 11.5084 | 12.46 | 11.16 | 10 | [5 empty] | 11.4013
| 1 | 5 | 12.2 | 18.1089 | 14.62 | 11.0103 | 12.9 | 10.64 | 10.9 | < 100 | 9.83 | 13.1 | 11.56 | 11.3081 | 12.12 | 10.92 | 10.1 | [5 empty] | 24
| 1 | 6 | 11.5 | 18.1278 | 13.91 | 11.3851 | 12.7 | 10.99 | 11.5 | < 100 | 9.78 | 12.5 | 11.08 | 11.1264 | 12.48 | 11.39 | 11.2
| Mean |  | 11.55 | 18.2503 | 14.6183 | 11.2354 | 12.5667 | 10.7767 | 11.15 |  | 9.7567 | 12.6333 | 11.6133 | 11.2901 | 12.4083 | 11.285 | 10.55
| Median |  | 11.5 | 18.2176 | 14.505 | 11.2438 | 12.6 | 10.79 | 11.15 |  | 9.755 | 12.75 | 11.64 | 11.3385 | 12.47 | 11.15 | 10.5
| Std Dev. |  | 0.35071 | 0.14654 | 0.60446 | 0.18588 | 0.36697 | 0.19002 | 0.20736 |  | 0.078909 | 0.55737 | 0.29575 | 0.17107 | 0.48783 | 0.4053 | 0.49295
| Rel.Std.Dev. |  | 3.036% | 0.803% | 4.135% | 1.654% | 2.92% | 1.763% | 1.86% |  | 0.8088% | 4.412% | 2.547% | 1.515% | 3.931% | 3.592% | 4.673%
| PDM3 |  | 1.304% | 60.07% | 28.22% | -1.455% | 10.22% | -5.479% | -2.204% |  | -14.42% | 10.81% | 1.86% | -0.9752% | 8.833% | -1.02% | -7.467%
| Z-Score (Absolute) |  | 0 | 5.26 | 2.41 | 0.25 | 0.8 | 0.61 | 0.31 | 30.21 | 1.41 | 0.85 | 0.05 | 0.2 | 0.67 | 0.21 | 0.79
| Analytical results for Ge in OREAS 466 (Certified Value 0.42 ppm) | [21 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.20 | 2.22 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [7 empty] | 2
2
| 1 | 1 | 1.9 | 1.02 | 0.4375 | 0.6026 | < 0.1 | 0.42 | 0.41 | 0.35 | 0.38 | 0.38 | 1.6 | 1.6 | 5.3 | [7 empty] | 1
| 1 | 2 | 1.9 | 1.05 | 0.4335 | 0.58788 | < 0.1 | 0.43 | 0.37 | 0.31 | 0.4 | 0.32 | 1.7 | 1.7 | 5.01 | [7 empty] | 5
| 1 | 3 | 1.9 | 1.08 | 0.46569 | 0.60787 | < 0.1 | 0.43 | 0.44 | 0.37 | 0.33 | 0.46 | 1.6 | 1.6 | 4.66 | [7 empty] | 16
| 1 | 4 | 2 | 1.29 | 0.43617 | 0.58549 | < 0.1 | 0.43 | 0.31 | 0.31 | 0.39 | 0.42 | 1.7 | 1.6 | 4.99 | [7 empty] | 0.42109
| 1 | 5 | 1.8 | 1.1 | 0.4256 | 0.58863 | < 0.1 | 0.45 | 0.42 | 0.34 | 0.34 | 0.37 | 1.7 | 1.7 | 4.92 | [7 empty] | 25
| 1 | 6 | 2 | 0.95 | 0.45064 | 0.59435 | < 0.1 | 0.41 | 0.4 | 0.32 | 0.37 | 0.39 | 1.7 | 1.6 | 5.2
| Mean |  | 1.9167 | 1.0817 | 0.44152 | 0.59447 |  | 0.42833 | 0.39167 | 0.33333 | 0.36833 | 0.39 | 1.6667 | 1.6333 | 5.0133
| Median |  | 1.9 | 1.065 | 0.43683 | 0.59149 |  | 0.43 | 0.405 | 0.33 | 0.375 | 0.385 | 1.7 | 1.6 | 5
| Std Dev. |  | 0.075277 | 0.11479 | 0.014352 | 0.0089878 |  | 0.013292 | 0.046224 | 0.024221 | 0.027869 | 0.047329 | 0.05164 | 0.05164 | 0.22411
| Rel.Std.Dev. |  | 3.928% | 10.61% | 3.251% | 1.512% |  | 3.103% | 11.8% | 7.266% | 7.566% | 12.14% | 3.098% | 3.162% | 4.47%
| PDM3 |  | 355.2% | 156.9% | 4.85% | 41.17% |  | 1.719% | -6.988% | -20.84% | -12.53% | -7.384% | 295.8% | 287.9% | 1091%
| Z-Score (Absolute) |  | NA | 7.65 | 0.36 | 2.1 | 4.1 | 0.21 | 0.21 | 0.87 | 0.47 | 0.23 | NA | NA | 52.46
Results from laboratories 1, 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 466 (Certified Value 4.7 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 4.89 | 4.28 | 3.7529 | 4.9948 | 4 | 4.41 | 4.6 | 5 | 4.8 | 4.9 | 5.06 | 4.8 | 4.5 | 6.8259 | 4.87 | 5.13 | 4.4 |  |  |  | 1
| 1 | 2 | 4.89 | 4.36 | 3.8208 | 5.0589 | 3.2 | 4.39 | 4.6 | 4.7 | 4.9 | 4.7 | 5.11 | 4.7 | 4.5 | 7.5565 | 4.65 | 4.96 | 4.7 |  |  |  | 0
| 1 | 3 | 4.87 | 4.11 | 3.2771 | 5.0471 | 4.3 | 4.41 | 4.4 | 4.7 | 4.7 | 4.8 | 4.42 | 4.8 | 4.6 | 7.5545 | 4.82 | 4.74 | 4.2 |  |  |  | 16
| 1 | 4 | 4.92 | 3.95 | 3.5867 | 5.1205 | 4.7 | 4.35 | 4.5 | 4.8 | 4.8 | 4.6 | 4.8 | 4.8 | 4.5 | 8.0137 | 4.86 | 4.73 | 4.4 |  |  |  | 4.6995
| 1 | 5 | 4.85 | 4.1 | 3.4961 | 5.0754 | 3.9 | 4.32 | 4.6 | 5.2 | 4.9 | 4.9 | 4.97 | 4.8 | 4.6 | 7.3953 | 4.69 | 4.74 | 4.3 |  |  |  | 26
| 1 | 6 | 4.99 | 4.08 | 3.33 | 5.184 | 4 | 4.48 | 4.7 | 5 | 4.7 | 4.7 | 4.72 | 4.8 | 4.7 | 6.9452 | 4.62 | 5.14 | 4.3
| Mean |  | 4.9017 | 4.1467 | 3.5439 | 5.0801 | 4.0167 | 4.3933 | 4.5667 | 4.9 | 4.8 | 4.7667 | 4.8467 | 4.7833 | 4.5667 | 7.3818 | 4.7517 | 4.9067 | 4.3833
| Median |  | 4.89 | 4.105 | 3.5414 | 5.0671 | 4 | 4.4 | 4.6 | 4.9 | 4.8 | 4.75 | 4.885 | 4.8 | 4.55 | 7.4749 | 4.755 | 4.85 | 4.35
| Std Dev. |  | 0.04916 | 0.14828 | 0.21972 | 0.065218 | 0.49565 | 0.055377 | 0.10328 | 0.2 | 0.089443 | 0.12111 | 0.25703 | 0.040825 | 0.08165 | 0.43805 | 0.11125 | 0.19694 | 0.17224
| Rel.Std.Dev. |  | 1.003% | 3.576% | 6.2% | 1.284% | 12.34% | 1.26% | 2.262% | 4.082% | 1.863% | 2.541% | 5.303% | 0.8535% | 1.788% | 5.934% | 2.341% | 4.014% | 3.929%
| PDM3 |  | 4.301% | -11.76% | -24.59% | 8.098% | -14.53% | -6.516% | -2.827% | 4.266% | 2.138% | 1.429% | 3.131% | 1.783% | -2.827% | 57.08% | 1.109% | 4.408% | -6.728%
| Z-Score (Absolute) |  | 0.46 | 2.09 | 4.12 | 1.06 | 2.53 | 1.26 | 0.67 | 0.45 | 0.11 | 0 | 0.27 | 0.06 | 0.67 | 8.82 | 0.05 | 0.47 | 1.29
| Analytical results for Hg in OREAS 466 (Indicative Value 0.47 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.02: Lab |  |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06 |  |  |  | ppm
|  |  | 2.01: 4A*OES/MS; 2.02: 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.01: 0.47; 2.02: < 1 |  |  |  | 1
| 1 | 2 | 2.01: 0.51; 2.02: < 1 |  |  |  | 2
| 1 | 3 | 2.01: 0.46; 2.02: < 1 |  |  |  | 16
| 1 | 4 | 2.01: 0.43; 2.02: < 1 |  |  |  | 0.46833
| 1 | 5 | 2.01: 0.47; 2.02: < 1 |  |  |  | 8
| 1 | 6 | 2.01: 0.47; 2.02: < 1
| Mean |  | 2.01: 0.46833
| Median |  | 2.01: 0.47
| Std Dev. |  | 2.01: 0.025626
| Rel.Std.Dev. |  | 2.01: 5.472%
| PDM3 |  | 2.01: 6.661e-14%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.67
| Analytical results for Ho in OREAS 466 (Certified Value 0.87 ppm) | [21 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 0.88 | 0.93266 | 1.073 | 0.84535 | 1 | 0.81 | 0.8 | 0.82 | 0.9 | 0.86 | 1.288 | 0.76 | 0.89 | 0.76 | [6 empty] | 1
| 1 | 2 | 0.9 | 0.9216 | 1.138 | 0.83969 | 1 | 0.72 | 0.8 | 0.82 | 0.9 | 0.856 | 1.2659 | 0.8 | 0.89 | 0.73 | [6 empty] | 0
| 1 | 3 | 0.9 | 0.96676 | 1.176 | 0.86188 | 1 | 0.77 | 0.8 | 0.86 | 0.8 | 0.86 | 1.2721 | 0.72 | 1.01 | 0.71 | [6 empty] | 16
| 1 | 4 | 0.91 | 0.92897 | 0.986 | 0.85072 | 1 | 0.76 | 0.8 | 0.84 | 0.9 | 0.85 | 1.3003 | 0.77 | 0.89 | 0.7 | [6 empty] | 0.86995
| 1 | 5 | 0.9 | 0.9271 | 0.996 | 0.82882 | 1 | 0.75 | 0.8 | 0.86 | 0.9 | 0.85 | 1.2918 | 0.74 | 0.89 | 0.69 | [6 empty] | 27
| 1 | 6 | 0.95 | 0.95693 | 1.191 | 0.87715 | 0.9 | 0.77 | 0.9 | 0.83 | 0.9 | 0.86 | 1.249 | 0.76 | 0.91 | 0.77
| Mean |  | 0.90667 | 0.939 | 1.0933 | 0.8506 | 0.98333 | 0.76333 | 0.81667 | 0.83833 | 0.88333 | 0.856 | 1.2779 | 0.75833 | 0.91333 | 0.72667
| Median |  | 0.9 | 0.93082 | 1.1055 | 0.84803 | 1 | 0.765 | 0.8 | 0.835 | 0.9 | 0.858 | 1.2801 | 0.76 | 0.89 | 0.72
| Std Dev. |  | 0.023381 | 0.018314 | 0.089191 | 0.017047 | 0.040825 | 0.029439 | 0.040825 | 0.018348 | 0.040825 | 0.004899 | 0.01904 | 0.027142 | 0.048028 | 0.03266
| Rel.Std.Dev. |  | 2.579% | 1.95% | 8.158% | 2.004% | 4.152% | 3.857% | 4.999% | 2.189% | 4.622% | 0.5723% | 1.49% | 3.579% | 5.259% | 4.494%
| PDM3 |  | 4.22% | 7.937% | 25.68% | -2.225% | 13.03% | -12.26% | -6.125% | -3.635% | 1.538% | -1.604% | 46.89% | -12.83% | 4.986% | -16.47%
| Z-Score (Absolute) |  | 0.41 | 0.76 | 2.47 | 0.21 | 1.25 | 1.17 | 0.58 | 0.35 | 0.15 | 0.15 | 4.5 | 1.23 | 0.48 | 1.58
| Analytical results for In in OREAS 466 (Certified Value 0.12 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.11 | 0.211 | 0.21515 | 0.1749 | 0.1 | 0.13 | 0.11 | 0.123 | 0.119 | 0.107 | 0.2 | 0.13 | 0.12 | 0.11456 | 0.094 | 0.12 | 0.12 | 0.203 |  |  | 1
| 1 | 2 | 0.13 | 0.203 | 0.2226 | 0.17879 | < 0.1 | 0.12 | 0.121 | 0.12 | 0.124 | 0.118 | 0.18 | 0.13 | 0.12 | 0.11972 | 0.093 | 0.15 | 0.12 | 0.227 |  |  | 6
| 1 | 3 | 0.12 | 0.208 | 0.24608 | 0.17821 | 0.1 | 0.12 | 0.12 | 0.125 | 0.127 | 0.129 | 0.17 | 0.14 | 0.13 | 0.11784 | 0.092 | 0.13 | 0.12 | 0.198 |  |  | 16
| 1 | 4 | 0.12 | 0.205 | 0.2343 | 0.17371 | 0.1 | 0.13 | 0.12 | 0.122 | 0.129 | 0.134 | 0.19 | 0.13 | 0.12 | 0.12286 | 0.098 | 0.13 | 0.11 | 0.248 |  |  | 0.12086
| 1 | 5 | 0.11 | 0.205 | 0.2511 | 0.16939 | 0.1 | 0.12 | 0.119 | 0.123 | 0.129 | 0.136 | 0.21 | 0.13 | 0.13 | 0.11841 | 0.089 | 0.13 | 0.11 | 0.219 |  |  | 28
| 1 | 6 | 0.13 | 0.183 | 0.23859 | 0.16989 | 0.1 | 0.13 | 0.114 | 0.129 | 0.106 | 0.132 | 0.2 | 0.13 | 0.13 | 0.11865 | 0.088 | 0.13 | 0.12 | 0.227
| Mean |  | 0.12 | 0.2025 | 0.23464 | 0.17415 | 0.1 | 0.125 | 0.11733 | 0.12367 | 0.12233 | 0.126 | 0.19167 | 0.13167 | 0.125 | 0.11867 | 0.092333 | 0.13167 | 0.11667 | 0.22033
| Median |  | 0.12 | 0.205 | 0.23644 | 0.17431 | 0.1 | 0.125 | 0.1195 | 0.123 | 0.1255 | 0.1305 | 0.195 | 0.13 | 0.125 | 0.11853 | 0.0925 | 0.13 | 0.12 | 0.223
| Std Dev. |  | 0.0089443 | 0.0099549 | 0.013728 | 0.0039883 | 0 | 0.0054772 | 0.0043665 | 0.0030768 | 0.0088468 | 0.011261 | 0.01472 | 0.0040825 | 0.0054772 | 0.0026947 | 0.0036148 | 0.0098319 | 0.005164 | 0.018195
| Rel.Std.Dev. |  | 7.454% | 4.916% | 5.851% | 2.29% | 0% | 4.382% | 3.721% | 2.488% | 7.232% | 8.937% | 7.68% | 3.101% | 4.382% | 2.271% | 3.915% | 7.467% | 4.426% | 8.258%
| PDM3 |  | -0.7129% | 67.55% | 94.14% | 44.09% | -17.26% | 3.424% | -2.919% | 2.321% | 1.218% | 4.251% | 58.58% | 8.94% | 3.424% | -1.811% | -23.6% | 8.94% | -3.471% | 82.3%
| Z-Score (Absolute) |  | 0.51 | 7.84 | 11.09 | 4.97 | NA | 0 | 0.78 | 0.13 | 0.27 | 0.1 | 6.74 | 0.67 | 0 | 0.64 | 3.3 | 0.67 | 0.84 | 9.64
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 466 (Certified Value 2.39 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 2.4271 | 2.3981 | 2.3428 | 1.571 | 2.3339 | 2.4 | 2.46 | 2.24 | 2.44 | 2.35 | 2.36 | 2.47 | 2.38 | 2.34 | 2.63 | 2.4216 | 2.44 | 2.44 |  |  | 1
| 1 | 2 | 2.4195 | 2.4364 | 2.3289 | 1.3927 | 2.3317 | 2.31 | 2.49 | 2.24 | 2.31 | 2.35 | 2.31 | 2.42 | 2.356 | 2.36 | 2.46 | 2.3578 | 2.42 | 2.43 |  |  | 0
| 1 | 3 | 2.4206 | 2.4522 | 2.3431 | 1.4312 | 2.3019 | 2.11 | 2.5 | 2.23 | 2.37 | 2.31 | 2.35 | 2.36 | 2.388 | 2.36 | 2.61 | 2.3726 | 2.45 | 2.4 |  |  | 16
| 1 | 4 | 2.3907 | 2.3854 | 2.3465 | 1.641 | 2.2817 | 2 | 2.5 | 2.22 | 2.53 | 2.37 | 2.33 | 2.33 | 2.376 | 2.35 | 2.36 | 2.4372 | 2.41 | 2.73 |  |  | 2.3934
| 1 | 5 | 2.4209 | 2.4274 | 2.339 | 1.4006 | 2.2574 | 2.14 | 2.48 | 2.27 | 2.45 | 2.35 | 2.37 | 2.45 | 2.38 | 2.34 | 2.67 | 2.4028 | 2.41 | 2.7 |  |  | 29
| 1 | 6 | 2.3985 | 2.431 | 2.3656 | 1.3288 | 2.3163 | 2.33 | 2.54 | 2.25 | 2.44 | 2.34 | 2.35 | 2.35 | 2.324 | 2.37 | 2.46 | 2.3482 | 2.45 | 1.91
| Mean |  | 2.4129 | 2.4217 | 2.3443 | 1.4609 | 2.3038 | 2.215 | 2.495 | 2.2417 | 2.4233 | 2.345 | 2.345 | 2.3967 | 2.3673 | 2.3533 | 2.5317 | 2.39 | 2.43 | 2.435
| Median |  | 2.42 | 2.4292 | 2.3429 | 1.4159 | 2.3091 | 2.225 | 2.495 | 2.24 | 2.44 | 2.35 | 2.35 | 2.39 | 2.378 | 2.355 | 2.535 | 2.3877 | 2.43 | 2.435
| Std Dev. |  | 0.01462 | 0.02506 | 0.012079 | 0.1193 | 0.029936 | 0.1545 | 0.026646 | 0.017224 | 0.075277 | 0.019748 | 0.021679 | 0.05785 | 0.023788 | 0.012111 | 0.12222 | 0.035992 | 0.018974 | 0.29467
| Rel.Std.Dev. |  | 0.6059% | 1.035% | 0.5152% | 8.166% | 1.299% | 6.975% | 1.068% | 0.7684% | 3.106% | 0.8422% | 0.9245% | 2.414% | 1.005% | 0.5146% | 4.827% | 1.506% | 0.7808% | 12.1%
| PDM3 |  | 0.8143% | 1.185% | -2.051% | -38.96% | -3.743% | -7.454% | 4.245% | -6.339% | 1.251% | -2.022% | -2.022% | 0.1367% | -1.089% | -1.674% | 5.777% | -0.1404% | 1.529% | 1.738%
| Z-Score (Absolute) |  | 0.53 | 0.66 | 0.53 | 14.11 | 1.15 | 2.52 | 1.79 | 2.11 | 0.69 | 0.52 | 0.52 | 0.28 | 0.17 | 0.39 | 2.35 | 0.17 | 0.79 | 0.87
| Analytical results for La in OREAS 466 (Certified Value 229 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 0
0
| 1 | 1 | 229.4 | 231.3 | 198.0883 | 265.8 | 211.3619 | 237 | 233.14 | 232 | 223 | 237 | 222 | 214 | 241.4 | 237.2 | 222 | 208.7922 | 252 | 222.89 | 260 | 240 | 1
| 1 | 2 | 215.1 | 232.6 | 197.3386 | 261.5 | 210.9311 | 226 | 225.62 | 226.1 | 217 | 225 | 215 | 217 | 237.1 | 234.8 | 229 | 195.3427 | 267 | 224.94 | 264 | 183 | 0
| 1 | 3 | 226.6 | 230.3 | 194.6357 | 263.2 | 213.4572 | 220 | 230.35 | 224.4 | 217 | 236 | 211 | 213 | 236.7 | 231 | 228 | 199.2639 | 245 | 223.61 | 257 | 226 | 16
| 1 | 4 | 221.7 | 223.6 | 201.1929 | 269.8 | 219.8986 | 230 | 224.18 | 220.1 | 212 | 246 | 217 | 217 | 247.2 | 232 | 232 | 197.8801 | 252 | 220.45 | 239 | 233 | 229.0675
| 1 | 5 | 229.1 | 230.3 | 200.6146 | 248.2 | 220.1946 | 228 | 236.97 | 223.8 | 224 | 240 | 217 | 218 | 249.8 | 232 | 219 | 213.8215 | 244 | 220.34 | 243 | 248 | 30
| 1 | 6 | 224.5 | 231.5 | 203.1288 | 263.8 | 219.3006 | 227 | 231.79 | 233.9 | 214 | 246 | 215 | 213 | 233.8 | 235 | 237 | 217.4239 | 255 | 227.55 | 270 | 243
| Mean |  | 224.4 | 229.9333 | 199.1665 | 262.05 | 215.8573 | 228 | 230.3417 | 226.7167 | 217.8333 | 238.3333 | 216.1667 | 215.3333 | 241 | 233.6667 | 227.8333 | 205.4207 | 252.5 | 223.2967 | 255.5 | 228.8333
| Median |  | 225.55 | 230.8 | 199.3514 | 263.5 | 216.3789 | 227.5 | 231.07 | 225.25 | 217 | 238.5 | 216 | 215.5 | 239.25 | 233.4 | 228.5 | 204.0281 | 252 | 223.25 | 258.5 | 236.5
| Std Dev. |  | 5.3985 | 3.2191 | 3.0638 | 7.3571 | 4.4099 | 5.5498 | 4.7775 | 5.245 | 4.7924 | 7.8145 | 3.6009 | 2.2509 | 6.3495 | 2.3754 | 6.5549 | 9.1908 | 8.3126 | 2.7532 | 12.112 | 23.7438
| Rel.Std.Dev. |  | 2.406% | 1.4% | 1.538% | 2.808% | 2.043% | 2.434% | 2.074% | 2.313% | 2.2% | 3.279% | 1.666% | 1.045% | 2.635% | 1.017% | 2.877% | 4.474% | 3.292% | 1.233% | 4.741% | 10.38%
| PDM3 |  | -2.038% | 0.378% | -13.05% | 14.4% | -5.767% | -0.466% | 0.5563% | -1.026% | -4.904% | 4.045% | -5.632% | -5.996% | 5.209% | 2.008% | -0.5388% | -10.32% | 10.23% | -2.519% | 11.54% | -0.1022%
| Z-Score (Absolute) |  | 0.23 | 0.13 | 1.89 | 2.25 | 0.79 | 0.01 | 0.16 | 0.08 | 0.66 | 0.69 | 0.77 | 0.83 | 0.86 | 0.38 | 0.01 | 1.48 | 1.62 | 0.3 | 1.81 | 0.06
| Analytical results for Li in OREAS 466 (Certified Value 43.9 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 43.7 | 44 | 52.8007 | 25.29 | 43.2238 | 44 | 44.1 | 41.4 | 43.7 | 43.6 | 43.4 | 46.6 | 47 | 43.4 | 39.392 | 42.4 | 50 | 50.2 |  |  | 1
| 1 | 2 | 44.8 | 46 | 52.1864 | 24.91 | 43.2066 | 44 | 44 | 41.1 | 41.8 | 43 | 42.6 | 44.9 | 47 | 44.5 | 37.0835 | 43 | 53 | 59.6 |  |  | 0
| 1 | 3 | 43.8 | 45 | 52.99 | 24.29 | 44.1688 | 43 | 44.5 | 41.2 | 42.3 | 42.2 | 43.2 | 43.6 | 48 | 43.3 | 36.7606 | 43.8 | 52 | 49 |  |  | 16
| 1 | 4 | 44.3 | 42 | 52.5959 | 24.45 | 43.6744 | 45 | 44.8 | 41 | 46.9 | 43.6 | 42.9 | 46.1 | 47.5 | 44.1 | 35.4933 | 42.5 | 47 | 43.6 |  |  | 43.8747
| 1 | 5 | 44.4 | 45 | 52.0714 | 24.04 | 44.0816 | 43 | 43.6 | 41.5 | 43.7 | 43.5 | 43.6 | 46.6 | 48.792 | 44.2 | 38.261 | 42.8 | 46 | 71.4 |  |  | 31
| 1 | 6 | 44.2 | 44 | 52.5307 | 23.62 | 43.8584 | 42 | 44.2 | 41.1 | 43.7 | 43 | 43 | 42.6 | 47.424 | 43.9 | 41.0316 | 41.6 | 51 | 50.6
| Mean |  | 44.2 | 44.3333 | 52.5292 | 24.4333 | 43.7023 | 43.5 | 44.2 | 41.2167 | 43.6833 | 43.15 | 43.1167 | 45.0667 | 47.6193 | 43.9 | 38.0037 | 42.6833 | 49.8333 | 54.0667
| Median |  | 44.25 | 44.5 | 52.5633 | 24.37 | 43.7664 | 43.5 | 44.15 | 41.15 | 43.7 | 43.25 | 43.1 | 45.5 | 47.462 | 44 | 37.6723 | 42.65 | 50.5 | 50.4
| Std Dev. |  | 0.40497 | 1.3663 | 0.35143 | 0.59969 | 0.41502 | 1.0488 | 0.41473 | 0.19408 | 1.7781 | 0.54314 | 0.36009 | 1.6753 | 0.68405 | 0.46904 | 1.9928 | 0.72778 | 2.7869 | 9.9323
| Rel.Std.Dev. |  | 0.9162% | 3.082% | 0.669% | 2.454% | 0.9497% | 2.411% | 0.9383% | 0.4709% | 4.07% | 1.259% | 0.8352% | 3.717% | 1.436% | 1.068% | 5.244% | 1.705% | 5.592% | 18.37%
| PDM3 |  | 0.7413% | 1.045% | 19.73% | -44.31% | -0.3931% | -0.8541% | 0.7413% | -6.058% | -0.4363% | -1.652% | -1.728% | 2.717% | 8.535% | 0.05758% | -13.38% | -2.715% | 13.58% | 23.23%
| Z-Score (Absolute) |  | 0.3 | 0.4 | 6.53 | 14.49 | 0.07 | 0.23 | 0.3 | 1.93 | 0.09 | 0.49 | 0.51 | 0.95 | 2.86 | 0.07 | 4.34 | 0.84 | 4.51 | 7.68
| Analytical results for Lu in OREAS 466 (Certified Value 0.28 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 0.27 | 0.32067 | 0.341 | 0.32569 | 0.4 | 0.2 | 0.3 | 0.27 | 0.3 | 0.28 | 0.4253 | 0.27 | 0.29 | 0.26 | [6 empty] | 1
| 1 | 2 | 0.29 | 0.3171 | 0.344 | 0.30043 | 0.4 | 0.21 | 0.3 | 0.28 | 0.3 | 0.274 | 0.43495 | 0.28 | 0.29 | 0.25 | [6 empty] | 0
| 1 | 3 | 0.27 | 0.31751 | 0.348 | 0.30151 | 0.3 | 0.21 | 0.3 | 0.28 | 0.3 | 0.278 | 0.4401 | 0.26 | 0.3 | 0.25 | [6 empty] | 16
| 1 | 4 | 0.29 | 0.29948 | 0.3 | 0.30241 | 0.4 | 0.24 | 0.3 | 0.28 | 0.3 | 0.276 | 0.45997 | 0.27 | 0.28 | 0.25 | [6 empty] | 0.28168
| 1 | 5 | 0.29 | 0.3444 | 0.341 | 0.28727 | 0.3 | 0.22 | 0.3 | 0.28 | 0.3 | 0.284 | 0.44406 | 0.26 | 0.28 | 0.24 | [6 empty] | 32
| 1 | 6 | 0.29 | 0.31096 | 0.406 | 0.30076 | 0.3 | 0.2 | 0.3 | 0.28 | 0.3 | 0.272 | 0.422 | 0.26 | 0.29 | 0.27
| Mean |  | 0.28333 | 0.31835 | 0.34667 | 0.30301 | 0.35 | 0.21333 | 0.3 | 0.27833 | 0.3 | 0.27733 | 0.43773 | 0.26667 | 0.28833 | 0.25333
| Median |  | 0.29 | 0.31731 | 0.3425 | 0.30114 | 0.35 | 0.21 | 0.3 | 0.28 | 0.3 | 0.277 | 0.43752 | 0.265 | 0.29 | 0.25
| Std Dev. |  | 0.010328 | 0.014813 | 0.033975 | 0.012463 | 0.054772 | 0.015055 | 0 | 0.0040825 | 0 | 0.0043205 | 0.013784 | 0.008165 | 0.0075277 | 0.010328
| Rel.Std.Dev. |  | 3.645% | 4.653% | 9.8% | 4.113% | 15.65% | 7.057% | 0% | 1.467% | 0% | 1.558% | 3.149% | 3.062% | 2.611% | 4.077%
| PDM3 |  | 0.5859% | 13.02% | 23.07% | 7.572% | 24.25% | -24.26% | 6.503% | -1.189% | 6.503% | -1.544% | 55.4% | -5.331% | 2.361% | -10.06%
| Z-Score (Absolute) |  | 0 | 1.2 | 2.17 | 0.67 | NA | 2.4 | NA | 0.17 | NA | 0.21 | 5.29 | 0.57 | 0.17 | 1.03
Results from laboratories 7, 9 and 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 466 (Certified Value 0.412 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.4388 | 0.3739 | 0.43604 | 0.40946 | 0.4412 | 0.52 | 0.39 | 0.4 | 0.43 | 0.41 | 0.41 | 0.45 | 0.3716 | 0.411 | 0.348 | 0.4629 | 0.38 | 0.42 |  |  | 1
| 1 | 2 | 0.4574 | 0.3764 | 0.43777 | 0.39251 | 0.4241 | 0.44 | 0.39 | 0.4 | 0.41 | 0.42 | 0.41 | 0.45 | 0.3656 | 0.42 | 0.365 | 0.439 | 0.38 | 0.38 |  |  | 0
| 1 | 3 | 0.4469 | 0.3709 | 0.43511 | 0.40267 | 0.43658 | 0.49 | 0.38 | 0.4 | 0.41 | 0.41 | 0.42 | 0.43 | 0.376 | 0.442 | 0.379 | 0.4503 | 0.38 | 0.41 |  |  | 16
| 1 | 4 | 0.4399 | 0.3647 | 0.43897 | 0.40321 | 0.43207 | 0.47 | 0.39 | 0.4 | 0.45 | 0.42 | 0.41 | 0.43 | 0.36 | 0.416 | 0.348 | 0.4723 | 0.38 | 0.43 |  |  | 0.41156
| 1 | 5 | 0.4488 | 0.3718 | 0.43691 | 0.39614 | 0.42708 | 0.53 | 0.39 | 0.4 | 0.42 | 0.42 | 0.41 | 0.43 | 0.3612 | 0.429 | 0.342 | 0.434 | 0.38 | 0.41 |  |  | 33
| 1 | 6 | 0.4468 | 0.3748 | 0.43253 | 0.38909 | 0.44434 | 0.44 | 0.39 | 0.41 | 0.42 | 0.41 | 0.41 | 0.42 | 0.3672 | 0.411 | 0.356 | 0.4244 | 0.38 | 0.43
| Mean |  | 0.44643 | 0.37208 | 0.43622 | 0.39885 | 0.43423 | 0.48167 | 0.38833 | 0.40167 | 0.42333 | 0.415 | 0.41167 | 0.435 | 0.36693 | 0.4215 | 0.35633 | 0.44715 | 0.38 | 0.41333
| Median |  | 0.44685 | 0.37285 | 0.43648 | 0.3994 | 0.43433 | 0.48 | 0.39 | 0.4 | 0.42 | 0.415 | 0.41 | 0.43 | 0.3664 | 0.418 | 0.352 | 0.44465 | 0.38 | 0.415
| Std Dev. |  | 0.0067405 | 0.0041306 | 0.0022486 | 0.0076041 | 0.0079333 | 0.038687 | 0.0040825 | 0.0040825 | 0.015055 | 0.0054772 | 0.0040825 | 0.012247 | 0.0061118 | 0.012079 | 0.013663 | 0.018152 | 6.0809e-17 | 0.018619
| Rel.Std.Dev. |  | 1.51% | 1.11% | 0.5155% | 1.907% | 1.827% | 8.032% | 1.051% | 1.016% | 3.556% | 1.32% | 0.9917% | 2.816% | 1.666% | 2.866% | 3.834% | 4.059% | 1.6e-14% | 4.505%
| PDM3 |  | 8.474% | -9.592% | 5.993% | -3.089% | 5.508% | 17.03% | -5.643% | -2.404% | 2.861% | 0.836% | 0.02612% | 5.696% | -10.84% | 2.415% | -13.42% | 8.648% | -7.668% | 0.4311%
| Z-Score (Absolute) |  | 1.01 | 1.32 | 0.69 | 0.48 | 0.63 | 2.12 | 0.81 | 0.39 | 0.29 | 0.03 | 0.08 | 0.66 | 1.49 | 0.23 | 1.82 | 1.04 | 1.07 | 0.03
| Analytical results for Mn in OREAS 466 (Certified Value 0.009 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0091 | 0.0084 | < 0.01 | 0.023998 | 0.0087489 | 0.0102 | 0.0097 | 0.0087 | 0.0094 | 0.0091 | 0.0089 | 0.0104 | 0.0089 | 0.0089 | 0.0099 | 0.0095 | 0.0084 | 0.0113 |  |  | 1
| 1 | 2 | 0.0091 | 0.0086 | < 0.01 | 0.019888 | 0.0087359 | 0.0102 | 0.0095 | 0.0087 | 0.009 | 0.009 | 0.0088 | 0.0101 | 0.0088 | 0.0089 | 0.0102 | 0.0091 | 0.0084 | 0.0117 |  |  | 0
| 1 | 3 | 0.0091 | 0.0085 | < 0.01 | 0.024301 | 0.0088177 | 0.0093 | 0.0095 | 0.0087 | 0.0091 | 0.009 | 0.009 | 0.0099 | 0.009 | 0.0089 | 0.0103 | 0.0091 | 0.0083 | 0.0111 |  |  | 16
| 1 | 4 | 0.0088 | 0.0084 | < 0.01 | 0.022368 | 0.0086632 | 0.0101 | 0.0095 | 0.0088 | 0.0099 | 0.0093 | 0.0089 | 0.0098 | 0.0086 | 0.009 | 0.0095 | 0.0094 | 0.0085 | 0.012 |  |  | 0.0091496
| 1 | 5 | 0.009 | 0.0086 | < 0.01 | 0.021026 | 0.0090151 | 0.0099 | 0.0098 | 0.0087 | 0.0093 | 0.0091 | 0.0089 | 0.0101 | 0.0088 | 0.0089 | 0.0101 | 0.0092 | 0.0085 | 0.0116 |  |  | 34
| 1 | 6 | 0.009 | 0.0089 | < 0.01 | 0.021758 | 0.0088823 | 0.0093 | 0.0096 | 0.0088 | 0.0093 | 0.009 | 0.009 | 0.0097 | 0.0087 | 0.0089 | 0.0099 | 0.0091 | 0.0084 | 0.0117
| Mean |  | 0.0090167 | 0.0085667 |  | 0.022223 | 0.0088105 | 0.0098333 | 0.0096 | 0.0087333 | 0.0093333 | 0.0090833 | 0.0089167 | 0.01 | 0.0088 | 0.0089167 | 0.0099833 | 0.0092333 | 0.0084167 | 0.011567
| Median |  | 0.00905 | 0.00855 |  | 0.022063 | 0.0087833 | 0.01 | 0.00955 | 0.0087 | 0.0093 | 0.00905 | 0.0089 | 0.01 | 0.0088 | 0.0089 | 0.01 | 0.00915 | 0.0084 | 0.01165
| Std Dev. |  | 0.0001169 | 0.00018619 |  | 0.0017084 | 0.000125 | 0.0004274 | 0.00012649 | 5.164e-05 | 0.00031411 | 0.0001169 | 7.5277e-05 | 0.00025298 | 0.00014142 | 4.0825e-05 | 0.00028577 | 0.00017512 | 7.5277e-05 | 0.00032042
| Rel.Std.Dev. |  | 1.297% | 2.173% |  | 7.687% | 1.419% | 4.346% | 1.318% | 0.5913% | 3.365% | 1.287% | 0.8442% | 2.53% | 1.607% | 0.4578% | 2.863% | 1.897% | 0.8944% | 2.77%
| PDM3 |  | -1.453% | -6.371% |  | 142.9% | -3.706% | 7.473% | 4.923% | -4.549% | 2.008% | -0.7242% | -2.546% | 9.294% | -3.821% | -2.546% | 9.112% | 0.9152% | -8.01% | 26.42%
| Z-Score (Absolute) |  | 0.06 | 0.81 | 6.83 | 22.21 | 0.4 | 1.32 | 0.93 | 0.53 | 0.48 | 0.06 | 0.22 | 1.6 | 0.42 | 0.22 | 1.57 | 0.31 | 1.07 | 4.24
| Analytical results for Mo in OREAS 466 (Certified Value 6.42 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 6.6 | 5.9 | 6.6084 | 3.517 | 7.0217 | < 1 | 6.28 | 6.2 | 5.97 | 6.54 | 6.27 | 6.12 | 7 | 6 | 6.2263 | 7.72 | 6.7 | 6.93 | 6.44 |  | 1
| 1 | 2 | 6.6 | 6.1 | 7.2203 | 3.9779 | 6.6432 | < 1 | 6.21 | 6.29 | 5.76 | 6.26 | 6.09 | 6.18 | 7.2 | 6 | 6.3961 | 7.25 | 6.5 | 6.99 | 6.8 |  | 0
| 1 | 3 | 6.5 | 5.9 | 6.817 | 5.894 | 7.1512 | 1 | 6.16 | 6.18 | 5.89 | 6.24 | 6.15 | 5.92 | 7.1 | 6.1 | 6.3304 | 7.52 | 6.7 | 6.8 | 6.54 |  | 16
| 1 | 4 | 6.6 | 6.1 | 7.7836 | 4.5357 | 6.6774 | 2 | 6.11 | 6.16 | 6.27 | 6.31 | 6.06 | 6.15 | 7 | 6.2 | 6.4601 | 7.99 | 6.9 | 6.48 | 6.8 |  | 6.4168
| 1 | 5 | 6.8 | 6.1 | 7.6118 | 4.0506 | 6.7852 | 1 | 6.22 | 6.23 | 6.15 | 6.36 | 6.29 | 6.39 | 7.1 | 6 | 6.2663 | 7.47 | 6.7 | 6.66 | 6.53 |  | 35
| 1 | 6 | 6.5 | 6.2 | 8.3837 | 4.9602 | 7.1439 | 2 | 6.26 | 6.11 | 6.27 | 6.28 | 6.03 | 5.94 | 7 | 6.3 | 6.3769 | 8.04 | 6.7 | 7.15 | 6.51
| Mean |  | 6.6 | 6.05 | 7.4041 | 4.4893 | 6.9038 | 1.5 | 6.2067 | 6.195 | 6.0517 | 6.3317 | 6.1483 | 6.1167 | 7.0667 | 6.1 | 6.3427 | 7.665 | 6.7 | 6.835 | 6.6033
| Median |  | 6.6 | 6.1 | 7.4161 | 4.2932 | 6.9034 | 1.5 | 6.215 | 6.19 | 6.06 | 6.295 | 6.12 | 6.135 | 7.05 | 6.05 | 6.3537 | 7.62 | 6.7 | 6.865 | 6.535
| Std Dev. |  | 0.10954 | 0.12247 | 0.65702 | 0.84812 | 0.23064 | 0.57735 | 0.06314 | 0.061563 | 0.21113 | 0.11035 | 0.10962 | 0.17305 | 0.08165 | 0.12649 | 0.086427 | 0.31002 | 0.12649 | 0.24089 | 0.15629
| Rel.Std.Dev. |  | 1.66% | 2.024% | 8.874% | 18.89% | 3.341% | 38.49% | 1.017% | 0.9938% | 3.489% | 1.743% | 1.783% | 2.829% | 1.155% | 2.074% | 1.363% | 4.045% | 1.888% | 3.524% | 2.367%
| PDM3 |  | 2.856% | -5.716% | 15.39% | -30.04% | 7.589% | -76.62% | -3.274% | -3.456% | -5.69% | -1.326% | -4.183% | -4.677% | 10.13% | -4.936% | -1.155% | 19.45% | 4.414% | 6.518% | 2.908%
| Z-Score (Absolute) |  | 0.64 | 0.7 | 2.61 | 4.52 | 1.39 | NA | 0.32 | 0.35 | 0.7 | 0.01 | 0.46 | 0.54 | 1.78 | 0.58 | 0.01 | 3.25 | 0.89 | 1.22 | 0.65
Results from laboratory 7 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 466 (Certified Value 0.164 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.1715 | 0.1378 | 0.20236 | 0.23306 | 0.17359 | 0.21 | 0.16 | 0.16 | 0.29 | 0.16 | 0.17 | 0.214 | 0.1442 | 0.156 | 0.484 | 0.1714 | 0.17 | 0.16 |  |  | 1
| 1 | 2 | 0.1786 | 0.14 | 0.20455 | 0.22049 | 0.17225 | 0.2 | 0.161 | 0.16 | 0.27 | 0.16 | 0.17 | 0.215 | 0.1436 | 0.157 | 0.478 | 0.1616 | 0.17 | 0.16 |  |  | 7
| 1 | 3 | 0.173 | 0.1378 | 0.20767 | 0.2159 | 0.17417 | 0.19 | 0.162 | 0.16 | 0.28 | 0.16 | 0.17 | 0.206 | 0.14609 | 0.16 | 0.477 | 0.1632 | 0.16 | 0.15 |  |  | 16
| 1 | 4 | 0.1722 | 0.1357 | 0.20121 | 0.21462 | 0.17053 | 0.2 | 0.164 | 0.16 | 0.18 | 0.17 | 0.17 | 0.205 | 0.14522 | 0.16 | 0.454 | 0.1768 | 0.16 | 0.16 |  |  | 0.1638
| 1 | 5 | 0.1725 | 0.1379 | 0.20918 | 0.20077 | 0.17312 | 0.2 | 0.16 | 0.16 | 0.29 | 0.17 | 0.17 | 0.213 | 0.14957 | 0.158 | 0.452 | 0.1601 | 0.16 | 0.16 |  |  | 36
| 1 | 6 | 0.1736 | 0.1419 | 0.20924 | 0.21585 | 0.17576 | 0.19 | 0.162 | 0.16 | 0.29 | 0.16 | 0.17 | 0.2 | 0.1488 | 0.158 | 0.489 | 0.1598 | 0.17 | 0.16
| Mean |  | 0.17357 | 0.13852 | 0.2057 | 0.21678 | 0.17324 | 0.19833 | 0.1615 | 0.16 | 0.26667 | 0.16333 | 0.17 | 0.20883 | 0.14624 | 0.15817 | 0.47233 | 0.16548 | 0.165 | 0.15833
| Median |  | 0.17275 | 0.13785 | 0.20611 | 0.21587 | 0.17335 | 0.2 | 0.1615 | 0.16 | 0.285 | 0.16 | 0.17 | 0.2095 | 0.14565 | 0.158 | 0.4775 | 0.1624 | 0.165 | 0.16
| Std Dev. |  | 0.0025665 | 0.0021442 | 0.0035006 | 0.010407 | 0.0017687 | 0.0075277 | 0.0015166 | 0 | 0.043205 | 0.005164 | 0 | 0.006047 | 0.0024418 | 0.0016021 | 0.015603 | 0.006995 | 0.0054772 | 0.0040825
| Rel.Std.Dev. |  | 1.479% | 1.548% | 1.702% | 4.801% | 1.021% | 3.795% | 0.9391% | 0% | 16.2% | 3.162% | 0% | 2.896% | 1.67% | 1.013% | 3.303% | 4.227% | 3.32% | 2.578%
| PDM3 |  | 5.964% | -15.43% | 25.58% | 32.35% | 5.761% | 21.08% | -1.403% | -2.319% | 62.8% | -0.284% | 3.786% | 27.49% | -10.72% | -3.438% | 188.4% | 1.029% | 0.7335% | -3.337%
| Z-Score (Absolute) |  | 0.41 | 2.08 | 2.7 | 3.48 | 0.39 | 2.17 | 0.44 | 0.55 | 7.03 | 0.31 | 0.16 | 2.92 | 1.53 | 0.68 | 21.64 | 0.16 | 0.19 | 0.67
| Analytical results for Nb in OREAS 466 (Certified Value 109 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 110.21 | 94.51 | 119.1833 | 113.1169 | 109 | 129.92 | 105.24 | 107.5 | 111 | 110.5 | 112 | 96.2 | 117.6 | 98.3 | 112.4796 | 124 | 107.53 | 143 | 108 |  | 1
| 1 | 2 | 111.24 | 96.95 | 143.2167 | 110.8769 | 106 | 126.69 | 105.12 | 105 | 106.5 | 107 | 108.5 | 98.04 | 114 | 101 | 113.7728 | 111 | 103.59 | 144 | 120 |  | 0
| 1 | 3 | 111.69 | 94.79 | 109.1917 | 116.5269 | 105 | 128.21 | 103.51 | 104.5 | 108.5 | 105 | 109.5 | 91.15 | 111.2 | 101 | 112.4643 | 116 | 106.85 | 140 | 115 |  | 16
| 1 | 4 | 112.08 | 92.35 | 108.35 | 114.6951 | 105 | 128.16 | 103.46 | 104.5 | 115.5 | 110 | 109 | 95.2 | 116 | 100 | 116.721 | 124 | 107.53 | 132 | 127 |  | 109.4805
| 1 | 5 | 110.03 | 95.36 | 101.1167 | 114.5822 | 107 | 117.83 | 105.22 | 106.5 | 113.5 | 107.5 | 112.5 | 96.28 | 112 | 101 | 111.085 | 124 | 106.48 | 134 | 112 |  | 37
| 1 | 6 | 111.53 | 94.69 | 131.675 | 116.0871 | 104 | 127.87 | 104.81 | 104 | 113 | 105.5 | 106.5 | 95.6 | 115 | 101 | 113.7518 | 121 | 107.79 | 145 | 111
| Mean |  | 111.13 | 94.775 | 118.7889 | 114.3142 | 106 | 126.4467 | 104.56 | 105.3333 | 111.3333 | 107.5833 | 109.6667 | 95.4117 | 114.3 | 100.3833 | 113.3791 | 120 | 106.6283 | 139.6667 | 115.5
| Median |  | 111.385 | 94.74 | 114.1875 | 114.6387 | 105.5 | 128.015 | 104.965 | 104.75 | 112 | 107.25 | 109.25 | 95.9 | 114.5 | 101 | 113.1157 | 122.5 | 107.19 | 141.5 | 113.5
| Std Dev. |  | 0.83001 | 1.4853 | 15.9769 | 2.0739 | 1.7889 | 4.3461 | 0.84695 | 1.3663 | 3.3566 | 2.2675 | 2.2509 | 2.3033 | 2.4191 | 1.0962 | 1.9168 | 5.4037 | 1.567 | 5.465 | 6.9498
| Rel.Std.Dev. |  | 0.7469% | 1.567% | 13.45% | 1.814% | 1.688% | 3.437% | 0.81% | 1.297% | 3.015% | 2.108% | 2.053% | 2.414% | 2.116% | 1.092% | 1.691% | 4.503% | 1.47% | 3.913% | 6.017%
| PDM3 |  | 1.507% | -13.43% | 8.502% | 4.415% | -3.179% | 15.5% | -4.494% | -3.788% | 1.692% | -1.733% | 0.1701% | -12.85% | 4.402% | -8.309% | 3.561% | 9.609% | -2.605% | 27.57% | 5.498%
| Z-Score (Absolute) |  | 0 | 2.15 | 1.01 | 0.42 | 0.67 | 2.01 | 0.86 | 0.76 | 0.03 | 0.47 | 0.19 | 2.07 | 0.42 | 1.41 | 0.3 | 1.17 | 0.59 | 3.75 | 0.57
| Analytical results for Nd in OREAS 466 (Certified Value 152 ppm) | [21 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 150.96 | 159.05 | 178.537 | 161.6 | 138.6611 | 164 | 144.72 | 157.8 | 158 | 200 | 200 | 150.5 | 158.9 | 146 | 149 | 141.3359 | 161 | 147.51 | 214 |  | 1
| 1 | 2 | 150.01 | 158.52 | 176.42 | 165.5 | 139.3498 | 155 | 145.16 | 153.9 | 159 | 200 | 200 | 148.5 | 158 | 145 | 153 | 137.9184 | 171 | 146.61 | 215 |  | 0
| 1 | 3 | 152.34 | 158.22 | 179.9574 | 159.4 | 142.9001 | 151 | 144.88 | 156 | 158.5 | 100 | 300 | 152.5 | 141 | 146 | 150 | 141.863 | 173 | 148.46 | 205 |  | 16
| 1 | 4 | 151.31 | 152.93 | 177.8314 | 157.2 | 141.2715 | 157 | 144.7 | 153.6 | 154.5 | 100 | 200 | 144 | 155.4 | 147 | 152 | 139.3178 | 161 | 146.59 | 196 |  | 152.0785
| 1 | 5 | 156.08 | 158.22 | 179.63 | 151.6 | 137.346 | 156 | 144.34 | 154.5 | 160.5 | 100 | 200 | 154 | 162.1 | 147 | 148 | 143.4322 | 152 | 145.5 | 198 |  | 38
| 1 | 6 | 154.89 | 158.3 | 178.7937 | 175.6 | 143.4669 | 157 | 145.5 | 160.8 | 159.5 | 200 | 300 | 146.5 | 147.7 | 145 | 158 | 141.932 | 162 | 150.87 | 218
| Mean |  | 152.5983 | 157.54 | 178.5283 | 161.8167 | 140.4992 | 156.6667 | 144.8833 | 156.1 | 158.3333 | 150 | 233.3333 | 149.3333 | 153.85 | 146 | 151.6667 | 140.9666 | 163.3333 | 147.59 | 207.6667
| Median |  | 151.825 | 158.26 | 178.6654 | 160.5 | 140.3106 | 156.5 | 144.8 | 155.25 | 158.75 | 150 | 200 | 149.5 | 156.7 | 146 | 151 | 141.5994 | 161.5 | 147.06 | 209.5
| Std Dev. |  | 2.3868 | 2.2801 | 1.2851 | 8.1869 | 2.4417 | 4.2269 | 0.40307 | 2.7799 | 2.0656 | 54.7723 | 51.6398 | 3.7506 | 7.9515 | 0.89443 | 3.6148 | 1.9981 | 7.6594 | 1.8895 | 9.3524
| Rel.Std.Dev. |  | 1.564% | 1.447% | 0.7198% | 5.059% | 1.738% | 2.698% | 0.2782% | 1.781% | 1.305% | 36.51% | 22.13% | 2.512% | 5.168% | 0.6126% | 2.383% | 1.417% | 4.689% | 1.28% | 4.504%
| PDM3 |  | 0.3418% | 3.591% | 17.39% | 6.403% | -7.614% | 3.017% | -4.731% | 2.644% | 4.113% | -1.367% | 53.43% | -1.805% | 1.165% | -3.997% | -0.2708% | -7.307% | 7.401% | -2.951% | 36.55%
| Z-Score (Absolute) |  | 0.13 | 0.4 | 2.66 | 0.86 | 1.44 | 0.3 | 0.97 | 0.24 | 0.48 | NA | NA | 0.49 | 0 | 0.85 | 0.24 | 1.39 | 1.02 | 0.67 | 5.8
| Analytical results for Ni in OREAS 466 (Certified Value 19 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 1
2
| 1 | 1 | 18.4 | 19 | 20.0141 | 14.196 | 18.9934 | 21 | 18.3 | 19.6 | 17.2 | 19.6 | 18 | 18.4 | 21.54 | 18 | 17.8364 | 19.1 | 20.7 | 20 |  |  | 1
| 1 | 2 | 19.3 | 19.4 | 20.1272 | 12.9669 | 18.7064 | 21 | 18.3 | 19.4 | 17.2 | 19 | 18 | 18.3 | 21.8 | 19 | 18.5866 | 17.6 | 21 | 18 |  |  | 0
| 1 | 3 | 18.9 | 18.4 | 19.9382 | 16.3916 | 18.287 | 21 | 18.3 | 19.5 | 18 | 18.6 | 17.8 | 17.8 | 21.3 | 19 | 18.063 | 18.3 | 20.7 | 22 |  |  | 16
| 1 | 4 | 19.1 | 17.7 | 19.87 | 18.0617 | 18.9426 | 22 | 18.4 | 20.5 | 18.4 | 19.5 | 17.8 | 18.9 | 21.079 | 18 | 19.0844 | 18.2 | 19.5 | 20 |  |  | 19.0014
| 1 | 5 | 18.3 | 18.8 | 21.1176 | 17.3956 | 18.3843 | 20 | 18.4 | 19.5 | 18.3 | 19.1 | 18.4 | 18.5 | 21.4 | 19 | 17.7766 | 17.7 | 19.5 | 19 |  |  | 39
| 1 | 6 | 19.1 | 18.2 | 22.4582 | 14.9936 | 19.299 | 21 | 18.5 | 19.1 | 18 | 19.4 | 17.4 | 18.3 | 21.6 | 18 | 18.295 | 18 | 21 | 21
| Mean |  | 18.85 | 18.5833 | 20.5876 | 15.6676 | 18.7688 | 21 | 18.3667 | 19.6 | 17.85 | 19.2 | 17.9 | 18.3667 | 21.4532 | 18.5 | 18.2737 | 18.15 | 20.4 | 20
| Median |  | 19 | 18.6 | 20.0706 | 15.6926 | 18.8245 | 21 | 18.35 | 19.5 | 18 | 19.25 | 17.9 | 18.35 | 21.47 | 18.5 | 18.179 | 18.1 | 20.7 | 20
| Std Dev. |  | 0.40866 | 0.608 | 1.0254 | 1.9573 | 0.38617 | 0.63246 | 0.08165 | 0.47329 | 0.5282 | 0.37417 | 0.32863 | 0.3559 | 0.25127 | 0.54772 | 0.49767 | 0.53944 | 0.70993 | 1.4142
| Rel.Std.Dev. |  | 2.168% | 3.272% | 4.981% | 12.49% | 2.058% | 3.012% | 0.4446% | 2.415% | 2.959% | 1.949% | 1.836% | 1.938% | 1.171% | 2.961% | 2.723% | 2.972% | 3.48% | 7.071%
| PDM3 |  | -0.7968% | -2.2% | 8.348% | -17.55% | -1.224% | 10.52% | -3.34% | 3.15% | -6.06% | 1.045% | -5.796% | -3.34% | 12.9% | -2.639% | -3.83% | -4.481% | 7.36% | 5.255%
| Z-Score (Absolute) |  | 0.18 | 0.1 | 1.98 | 3.12 | 0.1 | 2.41 | 0.32 | 0.96 | 0.86 | 0.54 | 0.8 | 0.32 | 2.88 | 0.18 | 0.42 | 0.54 | 1.79 | 1.37
| Analytical results for P in OREAS 466 (Certified Value 0.055 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.0536 | 0.0579 | 0.05753 | 0.052306 | 0.056082 | 0.049 | 0.054 | 0.053 | 0.057 | 0.055 | 0.056 | 0.057 | 0.0521 | 0.055 | 0.0541 | 0.06 | 0.053 |  |  |  | 1
| 1 | 2 | 0.0535 | 0.0569 | 0.057812 | 0.051762 | 0.055887 | 0.047 | 0.053 | 0.052 | 0.054 | 0.054 | 0.054 | 0.057 | 0.0522 | 0.055 | 0.053 | 0.06 | 0.059 |  |  |  | 0
| 1 | 3 | 0.0536 | 0.0581 | 0.055784 | 0.050537 | 0.056284 | 0.048 | 0.055 | 0.052 | 0.055 | 0.053 | 0.056 | 0.055 | 0.0524 | 0.055 | 0.0527 | 0.06 | 0.052 |  |  |  | 16
| 1 | 4 | 0.053 | 0.0543 | 0.059611 | 0.052038 | 0.057609 | 0.05 | 0.054 | 0.051 | 0.059 | 0.055 | 0.055 | 0.055 | 0.0524 | 0.055 | 0.0543 | 0.06 | 0.056 |  |  |  | 0.054981
| 1 | 5 | 0.054 | 0.0578 | 0.056166 | 0.051167 | 0.055274 | 0.047 | 0.054 | 0.052 | 0.057 | 0.054 | 0.056 | 0.056 | 0.052 | 0.055 | 0.0568 | 0.06 | 0.054 |  |  |  | 40
| 1 | 6 | 0.0533 | 0.0564 | 0.059167 | 0.050538 | 0.055869 | 0.044 | 0.054 | 0.052 | 0.057 | 0.054 | 0.055 | 0.055 | 0.052 | 0.055 | 0.0554 | 0.06 | 0.054
| Mean |  | 0.0535 | 0.0569 | 0.057678 | 0.051391 | 0.056168 | 0.0475 | 0.054 | 0.052 | 0.0565 | 0.054167 | 0.055333 | 0.055833 | 0.052183 | 0.055 | 0.054383 | 0.06 | 0.054667
| Median |  | 0.05355 | 0.05735 | 0.057671 | 0.051464 | 0.055985 | 0.0475 | 0.054 | 0.052 | 0.057 | 0.054 | 0.0555 | 0.0555 | 0.05215 | 0.055 | 0.0542 | 0.06 | 0.054
| Std Dev. |  | 0.00033466 | 0.0014325 | 0.0015404 | 0.00076153 | 0.00078295 | 0.0020736 | 0.00063246 | 0.00063246 | 0.0017607 | 0.00075277 | 0.0008165 | 0.00098319 | 0.00018348 | 0 | 0.0015303 | 0 | 0.0025033
| Rel.Std.Dev. |  | 0.6255% | 2.518% | 2.671% | 1.482% | 1.394% | 4.366% | 1.171% | 1.216% | 3.116% | 1.39% | 1.476% | 1.761% | 0.3516% | 0% | 2.814% | 0% | 4.579%
| PDM3 |  | -2.695% | 3.489% | 4.905% | -6.53% | 2.157% | -13.61% | -1.785% | -5.423% | 2.762% | -1.482% | 0.6399% | 1.549% | -5.089% | 0.03365% | -1.088% | 9.128% | -0.5726%
| Z-Score (Absolute) |  | 0.52 | 1 | 1.35 | 1.47 | 0.67 | 3.22 | 0.3 | 1.2 | 0.82 | 0.22 | 0.3 | 0.52 | 1.12 | 0.15 | 0.13 | 2.4 | 0
| Analytical results for Pb in OREAS 466 (Certified Value 23 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 1
2
| 1 | 1 | 23.4 | 26 | 20.9214 | 19.34 | 21.9827 | 20 | 24.69 | 21.5 | 24 | 21.1 | 23.4 | 23.7 | 24.1 | 23 | 23.4614 | 21.4 | 25 | 22 |  |  | 1
| 1 | 2 | 23.8 | 29 | 22.9206 | 19.38 | 22.0052 | 22 | 24.12 | 21.9 | 22.7 | 21.1 | 23.4 | 23.67 | 23.8 | 23 | 23.2552 | 21 | 25 | 24 |  |  | 4
| 1 | 3 | 23.6 | 25 | 23.472 | 17.65 | 22.0127 | 17 | 24.54 | 21.5 | 23.2 | 20.4 | 23.9 | 21.21 | 25.1 | 23 | 23.7049 | 21.7 | 24 | 49 |  |  | 16
| 1 | 4 | 24 | 27 | 21.4322 | 23.79 | 22.5173 | 17 | 24.41 | 22.2 | 24.6 | 21.1 | 23.1 | 24.13 | 23.4 | 23 | 23.5602 | 21.7 | 23 | 25 |  |  | 22.988
| 1 | 5 | 23.9 | 28 | 19.7408 | 22.21 | 22.7737 | 20 | 23.77 | 22.3 | 24.2 | 20.8 | 24.2 | 24.19 | 24.9 | 23 | 23.2812 | 21.3 | 26 | 26 |  |  | 41
| 1 | 6 | 23.8 | 31 | 22.496 | 16.92 | 22.2389 | 21 | 24.7 | 22.2 | 24 | 20.5 | 23.4 | 22.66 | 24.6 | 23 | 22.4507 | 21.7 | 26 | 39
| Mean |  | 23.75 | 27.6667 | 21.8305 | 19.8817 | 22.2551 | 19.5 | 24.3717 | 21.9333 | 23.7833 | 20.8333 | 23.5667 | 23.26 | 24.3167 | 23 | 23.2856 | 21.4667 | 24.8333 | 30.8333
| Median |  | 23.8 | 27.5 | 21.9641 | 19.36 | 22.1258 | 20 | 24.475 | 22.05 | 24 | 20.95 | 23.4 | 23.685 | 24.35 | 23 | 23.3713 | 21.55 | 25 | 25.5
| Std Dev. |  | 0.21679 | 2.1602 | 1.3912 | 2.6455 | 0.32662 | 2.0736 | 0.3645 | 0.36148 | 0.69976 | 0.32042 | 0.40332 | 1.1442 | 0.66156 | 0 | 0.44275 | 0.28752 | 1.169 | 10.7595
| Rel.Std.Dev. |  | 0.9128% | 7.808% | 6.373% | 13.31% | 1.468% | 10.63% | 1.496% | 1.648% | 2.942% | 1.538% | 1.711% | 4.919% | 2.721% | 0% | 1.901% | 1.339% | 4.708% | 34.9%
| PDM3 |  | 3.315% | 20.35% | -5.035% | -13.51% | -3.188% | -15.17% | 6.019% | -4.588% | 3.46% | -9.373% | 2.517% | 1.183% | 5.78% | 0.05207% | 1.295% | -6.618% | 8.027% | 34.13%
| Z-Score (Absolute) |  | 0.26 | 2.43 | 0.8 | 1.88 | 0.56 | 2.09 | 0.61 | 0.74 | 0.28 | 1.35 | 0.16 | 0.01 | 0.58 | 0.15 | 0.01 | 1 | 0.86 | 4.18
| Analytical results for Pr in OREAS 466 (Certified Value 44.1 ppm) | [21 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 45.13 | 53.0207 | 51.37 | 44.8323 | 48.3 | 42.79 | 45.5 | < 100 | 44.8 | 45.3 | 45.8768 | 41.0393 | 43.29 | 43.13 | 64 | [5 empty] | 1
| 1 | 2 | 43.26 | 52.392 | 50.36 | 44.8537 | 42.9 | 42.65 | 43.6 | < 100 | 44.6 | 45.5 | 44 | 38.635 | 43.59 | 42.65 | 64.1 | [5 empty] | 0
| 1 | 3 | 44.82 | 52.9259 | 50.56 | 45.9699 | 43.2 | 42.71 | 43.5 | < 100 | 44 | 39.5 | 45.8768 | 40.1876 | 41.58 | 43.48 | 62.6 | [5 empty] | 16
| 1 | 4 | 42.73 | 52.8111 | 52.42 | 44.484 | 45 | 42.95 | 43.4 | < 100 | 43.7 | 44.9 | 44 | 39.5333 | 44.32 | 43.27 | 59.5 | [5 empty] | 44.1153
| 1 | 5 | 46.22 | 53.006 | 48.07 | 43.5013 | 45 | 42.79 | 43.9 | < 100 | 45.7 | 46.8 | 45.1185 | 42.0766 | 41.78 | 42.55 | 59.3 | [5 empty] | 42
| 1 | 6 | 44.97 | 53.4659 | 50.9 | 44.5531 | 45.4 | 42.62 | 45.1 | < 100 | 44.5 | 42.1 | 44.8 | 41.3242 | 43.82 | 44.33 | 66
| Mean |  | 44.5217 | 52.9369 | 50.6133 | 44.6991 | 44.9667 | 42.7517 | 44.1667 |  | 44.55 | 44.0167 | 44.9453 | 40.466 | 43.0633 | 43.235 | 62.5833
| Median |  | 44.895 | 52.9659 | 50.73 | 44.6927 | 45 | 42.75 | 43.75 |  | 44.55 | 45.1 | 44.9592 | 40.6135 | 43.44 | 43.2 | 63.3
| Std Dev. |  | 1.292 | 0.34772 | 1.446 | 0.79484 | 1.9336 | 0.11974 | 0.90259 |  | 0.69498 | 2.6999 | 0.84541 | 1.2617 | 1.125 | 0.64541 | 2.6933
| Rel.Std.Dev. |  | 2.902% | 0.6569% | 2.857% | 1.778% | 4.3% | 0.2801% | 2.044% |  | 1.56% | 6.134% | 1.881% | 3.118% | 2.612% | 1.493% | 4.303%
| PDM3 |  | 0.9212% | 20% | 14.73% | 1.323% | 1.93% | -3.091% | 0.1165% |  | 0.9854% | -0.2235% | 1.882% | -8.272% | -2.385% | -1.995% | 41.86%
| Z-Score (Absolute) |  | 0.01 | 4.3 | 3.11 | 0.08 | 0.21 | 0.92 | 0.2 | 2.79 | 0 | 0.27 | 0.2 | 2.09 | 0.76 | 0.67 | 9.25
| Analytical results for Rb in OREAS 466 (Certified Value 140 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 139.7 | 148.1 | 156.4653 | 143.3 | 144.1139 | 149 | 138.1 | 143 | 143.5 | 138.5 | 127 | 141.6 | 135.2 | 139 | 109.5683 | 146 | 139.89 | 154 | 123 |  | 1
| 1 | 2 | 138.4 | 150.6 | 154.61 | 138.5 | 142.321 | 142 | 136.7 | 142 | 141.5 | 137 | 130.5 | 141.3 | 136.7 | 139 | 66.81 | 147 | 140.12 | 152 | 128 |  | 0
| 1 | 3 | 136.7 | 147.6 | 156.6628 | 135.7 | 146.6144 | 133 | 137.1 | 141.5 | 148 | 133.5 | 128 | 127.4 | 135.7 | 138 | 66.0855 | 143 | 140.78 | 150 | 125 |  | 16
| 1 | 4 | 134.4 | 136.4 | 155.8469 | 145.1 | 146.108 | 131 | 136.4 | 140 | 148 | 136 | 127 | 135.6 | 135.6 | 139 | 53.7177 | 145 | 140.68 | 134 | 121 |  | 140.2327
| 1 | 5 | 138.6 | 148.2 | 159.86 | 135.7 | 143.8247 | 138 | 137.9 | 146.5 | 144.5 | 137.5 | 133 | 151.2 | 135.4 | 140 | 87.2379 | 142 | 142.39 | 139 | 121 |  | 43
| 1 | 6 | 140 | 142.8 | 158.2614 | 142.2 | 148.1665 | 140 | 139.7 | 140 | 151.5 | 138.5 | 126.5 | 131.5 | 135.2 | 138 | 124.4778 | 142 | 141.03 | 153 | 132
| Mean |  | 137.9667 | 145.6167 | 156.9511 | 140.0833 | 145.1914 | 138.8333 | 137.65 | 142.1667 | 146.1667 | 136.8333 | 128.6667 | 138.1 | 135.6333 | 138.8333 | 84.6495 | 144.1667 | 140.815 | 147 | 125
| Median |  | 138.5 | 147.85 | 156.5641 | 140.35 | 145.111 | 139 | 137.5 | 141.75 | 146.25 | 137.25 | 127.5 | 138.45 | 135.5 | 139 | 77.0239 | 144 | 140.73 | 151 | 124
| Std Dev. |  | 2.1002 | 5.1863 | 1.8546 | 4.0231 | 2.144 | 6.4936 | 1.2029 | 2.4221 | 3.656 | 1.8886 | 2.5626 | 8.4664 | 0.56095 | 0.75277 | 27.6897 | 2.137 | 0.88147 | 8.3905 | 4.3359
| Rel.Std.Dev. |  | 1.522% | 3.562% | 1.182% | 2.872% | 1.477% | 4.677% | 0.8739% | 1.704% | 2.501% | 1.38% | 1.992% | 6.131% | 0.4136% | 0.5422% | 32.71% | 1.482% | 0.626% | 5.708% | 3.469%
| PDM3 |  | -1.616% | 3.839% | 11.92% | -0.1065% | 3.536% | -0.9979% | -1.842% | 1.379% | 4.232% | -2.424% | -8.248% | -1.521% | -3.28% | -0.9979% | -39.64% | 2.805% | 0.4152% | 4.826% | -10.86%
| Z-Score (Absolute) |  | 0.18 | 1.37 | 3.67 | 0.25 | 1.29 | 0 | 0.24 | 0.67 | 1.48 | 0.4 | 2.06 | 0.15 | 0.65 | 0 | 10.96 | 1.08 | 0.4 | 1.65 | 2.8
| Analytical results for Re in OREAS 466 (Indicative Value 0.002 ppm) | [21 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.22 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | 0.003 | < 0.01 | < 0.001 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.05 | < 0.05 | 0.003 | [8 empty] | 1
| 1 | 2 | 0.004 | < 0.01 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.004 | < 0.05 | < 0.05 | 0.002 | [8 empty] | 1
| 1 | 3 | < 0.003 | < 0.01 | 0.001 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.003 | < 0.05 | < 0.05 | 0.002 | [8 empty] | 16
| 1 | 4 | 0.003 | < 0.01 | < 0.001 | < 0.002 | 0.003 | 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.05 | < 0.05 | 0.002 | [8 empty] | 0.0018958
| 1 | 5 | 0.003 | < 0.01 | < 0.001 | < 0.002 | 0.003 | < 0.002 | 0.002 | 0.003 | < 0.002 | < 0.05 | < 0.05 | 0.002 | [8 empty] | 7
| 1 | 6 | < 0.003 | < 0.01 | < 0.001 | < 0.002 | 0.003 | 0.002 | 0.003 | < 0.002 | 0.002 | < 0.05 | < 0.05 | 0.002
| Mean |  | 2.01: 0.00325; 2.04: 0.0015; 2.06: 0.0025; 2.07: 0.002; 2.09: 0.0022; 2.10: 0.0025; 2.11: 0.00275; 2.15: 0.0021667
| Median |  | 2.01: 0.003; 2.04: 0.0015; 2.06: 0.0025; 2.07: 0.002; 2.09: 0.002; 2.10: 0.0025; 2.11: 0.0025; 2.15: 0.002
| Std Dev. |  | 2.01: 0.0005; 2.04: 0.00070711; 2.06: 0.00054772; 2.07: 0; 2.09: 0.00044721; 2.10: 0.00070711; 2.11: 0.00095743; 2.15: 0.00040825
| Rel.Std.Dev. |  | 2.01: 15.38%; 2.04: 47.14%; 2.06: 21.91%; 2.07: 0%; 2.09: 20.33%; 2.10: 28.28%; 2.11: 34.82%; 2.15: 18.84%
| PDM3 |  | 2.01: 71.43%; 2.04: -20.88%; 2.06: 31.87%; 2.07: 5.495%; 2.09: 16.04%; 2.10: 31.87%; 2.11: 45.05%; 2.15: 14.29%
| Z-Score (Absolute) |  | 0.45 | 2.55 | 1.2 | 1.05 | 0.3 | 0.75 | 0.15 | 0.6 | 0 | 20.53 | 20.53 | 0
| Analytical results for S in OREAS 466 (Certified Value 0.012 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.19 | 2.21 | 2.22 | [5 empty] | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | 0.01 | 0.0115 | 0.016191 | 0.038836 | < 0.01 | < 0.04 | 0.01 | 0.01 | 0.01 | 0.03 | < 0.04 | 0.008 | 0.015 | < 0.01 | < 0.01 | [5 empty] | 1
| 1 | 2 | 0.02 | 0.0116 | 0.016504 | 0.033671 | < 0.01 | < 0.04 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.04 | 0.008 | 0.013 | < 0.01 | < 0.01 | [5 empty] | 6
| 1 | 3 |  | 0.0131 | 0.016611 | 0.037447 | 0.02 | < 0.04 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.04 | 0.008 | 0.013 | < 0.01 | < 0.01 | [5 empty] | 16
| 1 | 4 |  | 0.0104 | 0.016358 | 0.034422 | 0.02 | < 0.04 | 0.01 | 0.02 | 0.01 | 0.02 | < 0.04 | 0.008 | 0.015 | < 0.01 | < 0.01 | [5 empty] | 0.011701
| 1 | 5 | 0.01 | 0.0093 | 0.016139 | 0.035725 | < 0.01 | < 0.04 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.04 | 0.008 | 0.014 | < 0.01 | < 0.01 | [5 empty] | 44
| 1 | 6 | 0.01 | 0.0117 | 0.016254 | 0.033606 | < 0.01 | < 0.04 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.04 | 0.008 | 0.013 | < 0.01 | < 0.01
| Mean |  | 2.01: 0.0125; 2.02: 0.011267; 2.04: 0.016343; 2.05: 0.035618; 2.06: 0.02; 2.09: 0.01; 2.10: 0.011667; 2.11: 0.01; 2.12: 0.021667; 2.15: 0.008; 2.16: 0.013833
| Median |  | 2.01: 0.01; 2.02: 0.01155; 2.04: 0.016306; 2.05: 0.035074; 2.06: 0.02; 2.09: 0.01; 2.10: 0.01; 2.11: 0.01; 2.12: 0.02; 2.15: 0.008; 2.16: 0.0135
| Std Dev. |  | 2.01: 0.005; 2.02: 0.001291; 2.04: 0.00018447; 2.05: 0.002144; 2.06: 0; 2.09: 0; 2.10: 0.0040825; 2.11: 0; 2.12: 0.0040825; 2.15: 0; 2.16: 0.00098319
| Rel.Std.Dev. |  | 2.01: 40%; 2.02: 11.46%; 2.04: 1.129%; 2.05: 6.019%; 2.06: 0%; 2.09: 0%; 2.10: 34.99%; 2.11: 0%; 2.12: 18.84%; 2.15: 0%; 2.16: 7.107%
| PDM3 |  | 2.01: 6.827%; 2.02: -3.714%; 2.04: 39.67%; 2.05: 204.4%; 2.06: 70.92%; 2.09: -14.54%; 2.10: -0.2951%; 2.11: -14.54%; 2.12: 85.17%; 2.15: -31.63%; 2.16: 18.22%
| Z-Score (Absolute) |  | 0.15 | 0.07 | 0.86 | 4.4 | 0.31 | 1.53 | 0.31 | 0 | 0.31 | 1.84 | 1.53 | 0.67 | 0.4 | 1.23 | 1.23
| Analytical results for Sb in OREAS 466 (Certified Value 5.45 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 5.2 | 5.53 | 10.2594 |  | 5.1892 | 0.4 | 5.29 | 5.23 | 5.45 | 5.44 | 5.39 | 5.87 | 5.6 | 5.4 | 5.4273 | 5.66 | 3.52 | 9.85 |  |  | 1
| 1 | 2 | 5.4 | 5.6 | 9.4374 | 0.26018 | 5.2653 | 0.1 | 5.24 | 5.19 | 5.34 | 5.39 | 5.57 | 5.71 | 5.4 | 5.5 | 5.5277 | 5.79 | 3.44 | 70.3 |  |  | 7
| 1 | 3 | 5.6 | 5.48 | 8.9568 |  | 5.315 | 0.1 | 5.24 | 5.21 | 5.3 | 5.3 | 5.52 | 5.08 | 5.6 | 5.5 | 5.3304 | 5.46 | 3.08 | 6.27 |  |  | 16
| 1 | 4 | 5.5 | 5.43 | 9.9674 |  | 5.2756 | 0.2 | 5.34 | 4.98 | 5.7 | 5.55 | 5.64 | 5.45 | 5.5 | 5.4 | 5.4068 | 5.68 | 3.31 | 8.17 |  |  | 5.4451
| 1 | 5 | 6.1 | 5.44 | 9.2208 | 0.59638 | 5.2759 | < 0.1 | 5.18 | 5.24 | 6.07 | 5.44 | 5.84 | 5.91 | 5.6 | 5.5 | 5.294 | 5.65 | 3.27 | 6.06 |  |  | 45
| 1 | 6 | 5.7 | 5.47 | 10.0128 | 0.39703 | 5.4852 | 0.1 | 5.35 | 5.08 | 5.59 | 5.33 | 5.4 | 5.26 | 5.5 | 5.5 | 5.4027 | 5.56 | 3.51 | 8.15
| Mean |  | 5.5833 | 5.4917 | 9.6424 | 0.41786 | 5.301 | 0.18 | 5.2733 | 5.155 | 5.575 | 5.4083 | 5.56 | 5.5467 | 5.5333 | 5.4667 | 5.3981 | 5.6333 | 3.355 | 18.1333
| Median |  | 5.55 | 5.475 | 9.7024 | 0.39703 | 5.2757 | 0.1 | 5.265 | 5.2 | 5.52 | 5.415 | 5.545 | 5.58 | 5.55 | 5.5 | 5.4048 | 5.655 | 3.375 | 8.16
| Std Dev. |  | 0.30605 | 0.063692 | 0.51251 | 0.16907 | 0.099154 | 0.13038 | 0.065625 | 0.1033 | 0.28543 | 0.089759 | 0.16793 | 0.33839 | 0.08165 | 0.05164 | 0.081429 | 0.11237 | 0.16932 | 25.5946
| Rel.Std.Dev. |  | 5.481% | 1.16% | 5.315% | 40.46% | 1.87% | 72.44% | 1.244% | 2.004% | 5.12% | 1.66% | 3.02% | 6.101% | 1.476% | 0.9446% | 1.508% | 1.995% | 5.047% | 141.1%
| PDM3 |  | 2.54% | 0.8561% | 77.09% | -92.33% | -2.645% | -96.69% | -3.154% | -5.327% | 2.387% | -0.6744% | 2.111% | 1.866% | 1.621% | 0.3969% | -0.8614% | 3.458% | -38.38% | 233%
| Z-Score (Absolute) |  | 0.54 | 0.07 | 21.73 | 26.42 | 0.93 | 27.78 | 1.07 | 1.69 | 0.5 | 0.37 | 0.42 | 0.35 | 0.28 | 0.07 | 0.42 | 0.8 | 11.09 | 66.06
| Analytical results for Sc in OREAS 466 (Certified Value 18.8 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 18.5 | 17.85 | 18.7125 | 24.41 | 19.2519 | 21 | 20.59 | 17 | 18.4 | 17.8 | 18.4 | 17.7 | 21.8 | 19.41 | 18 | 14.9625 | 18.7 | 22.1 | 16 |  | 1
| 1 | 2 | 18.3 | 18.11 | 18.5643 | 23.71 | 18.8921 | 21 | 19.87 | 17 | 17.8 | 18.2 | 18.2 | 17.6 | 19.8 | 20.7207 | 18 | 13.0558 | 18.5 | 22.6 | 18.8 |  | 0
| 1 | 3 | 18.7 | 18.07 | 18.341 | 21.76 | 19.3051 | 20 | 20.39 | 16.9 | 18.2 | 18.9 | 17.6 | 17.5 | 20.6 | 19.8198 | 19 | 13.277 | 19 | 22.5 | 15.9 |  | 16
| 1 | 4 | 18.6 | 17.68 | 18.9665 | 21.54 | 19.2821 | 21 | 19.83 | 16.7 | 18 | 19.4 | 17.9 | 17.6 | 20.8 | 19.8198 | 18 | 12.5268 | 18.4 | 19.9 | 18.5 |  | 18.8402
| 1 | 5 | 18.6 | 17.94 | 18.7661 | 19.96 | 19.5391 | 21 | 20.03 | 16.8 | 18.8 | 18.4 | 18.2 | 17.6 | 20.6 | 20.7207 | 18 | 15.4576 | 18.9 | 19 | 16.5 |  | 46
| 1 | 6 | 18.6 | 17.88 | 18.849 | 23.63 | 19.2118 | 21 | 19.77 | 17.2 | 17.6 | 18.2 | 18.1 | 17.2 | 20 | 20.1 | 18 | 15.9395 | 18.8 | 23.2 | 16.3
| Mean |  | 18.55 | 17.9217 | 18.6999 | 22.5017 | 19.247 | 20.8333 | 20.08 | 16.9333 | 18.1333 | 18.4833 | 18.0667 | 17.5333 | 20.6 | 20.0985 | 18.1667 | 14.2032 | 18.7167 | 21.55 | 17
| Median |  | 18.6 | 17.91 | 18.7393 | 22.695 | 19.267 | 21 | 19.95 | 16.95 | 18.1 | 18.3 | 18.15 | 17.6 | 20.6 | 19.9599 | 18 | 14.1197 | 18.75 | 22.3 | 16.4
| Std Dev. |  | 0.13784 | 0.1569 | 0.22139 | 1.6917 | 0.2084 | 0.40825 | 0.33508 | 0.17512 | 0.43205 | 0.57417 | 0.28048 | 0.17512 | 0.70427 | 0.52984 | 0.40825 | 1.4247 | 0.23166 | 1.6885 | 1.2992
| Rel.Std.Dev. |  | 0.7431% | 0.8755% | 1.184% | 7.518% | 1.083% | 1.96% | 1.669% | 1.034% | 2.383% | 3.106% | 1.552% | 0.9988% | 3.419% | 2.636% | 2.247% | 10.03% | 1.238% | 7.835% | 7.643%
| PDM3 |  | -1.54% | -4.876% | -0.7447% | 19.43% | 2.159% | 10.58% | 6.58% | -10.12% | -3.752% | -1.894% | -4.106% | -6.937% | 9.341% | 6.679% | -3.575% | -24.61% | -0.6558% | 14.38% | -9.768%
| Z-Score (Absolute) |  | 0 | 0.42 | 0.1 | 2.62 | 0.46 | 1.51 | 1.01 | 1.07 | 0.28 | 0.04 | 0.32 | 0.67 | 1.36 | 1.03 | 0.25 | 2.88 | 0.11 | 1.99 | 1.03
| Analytical results for Se in OREAS 466 (Indicative Value 0.9 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 0.5 | 3.26 | 17.3558 | 2.9578 | < 0.1 | 0.4 | 1 | < 1 | 1 | 2 | 0.7 | NA | < 1 | 0.5 | < 2 | 4 |  |  |  |  | 1
| 1 | 2 | 0.5 | 4.71 | 17.7674 | 3.3181 | < 0.1 | 0.4 | 1 | < 1 | 1 | 2 | 0.6 | NA | < 1 | 0.6 | < 2 | 5 |  |  |  |  | 2
| 1 | 3 | 0.6 | 3.6 | 17.3904 | 2.9269 | < 0.1 | 0.5 | 1 | < 1 | 1 | 2 | 0.6 | NA | 1 | 0.5 | < 2 | 3 |  |  |  |  | 16
| 1 | 4 | 0.5 | 2.42 | 16.8046 | 2.8635 | < 0.1 | 0.6 | 1 | 1 | 1 | 2 | 0.4 | NA | 1 | 0.6 | < 2 | 5 |  |  |  |  | 0.90185
| 1 | 5 | 0.5 | 4.85 | 18.1824 | 3.2686 | < 0.1 | 0.5 | 1 | < 1 | 1 | 2 | 0.7 | NA | < 1 | 0.5 | < 2 | 5 |  |  |  |  | 8
| 1 | 6 | 0.5 | 2.84 | 16.4064 | 3.1355 | < 0.1 | 0.3 | 1 | < 1 | 1 | 2 | < 0.3 | NA | < 1 | 0.6 | < 2 | 4
| Mean |  | 0.51667 | 3.6133 | 17.3178 | 3.0784 |  | 0.45 | 1 | 1 | 1 | 2 | 0.6 |  | 1 | 0.55 |  | 4.3333
| Median |  | 0.5 | 3.43 | 17.3731 | 3.0467 |  | 0.45 | 1 | 1 | 1 | 2 | 0.6 |  | 1 | 0.55 |  | 4.5
| Std Dev. |  | 0.040825 | 0.98784 | 0.64038 | 0.19001 |  | 0.10488 | 0 |  | 0 | 0 | 0.12247 |  | 0 | 0.054772 |  | 0.8165
| Rel.Std.Dev. |  | 7.902% | 27.34% | 3.698% | 6.173% |  | 23.31% | 0% |  | 0% | 0% | 20.41% |  | 0% | 9.959% |  | 18.84%
| PDM3 |  | -42.71% | 300.7% | 1820% | 241.3% |  | -50.1% | 10.88% | 10.88% | 10.88% | 121.8% | -33.47% |  | 10.88% | -39.01% |  | 380.5%
| Z-Score (Absolute) |  | 0.67 | 3.65 | 22.77 | 2.9 | 1.33 | 0.77 | 0 | 0.58 | 0 | 1.4 | 0.66 | NA | 0.47 | 0.63 | 0 | 4.65
| Analytical results for Sm in OREAS 466 (Certified Value 20.8 ppm) | [21 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 21.5 | 22.0579 | 22.64 | 20.0822 | 18.1 | 22.62 | 21.3 | < 100 | 19.9 | 21.7 | 21.56 | 19.3944 | 21.2 | 20.63 | 24.3 | [5 empty] | 1
| 1 | 2 | 20.7 | 22.6975 | 20.53 | 20.4093 | 19.9 | 21.42 | 21 | < 100 | 20.5 | 21.3 | 21.44 | 19.3758 | 22.23 | 20.52 | 24.5 | [5 empty] | 0
| 1 | 3 | 20.6 | 22.0018 | 20.75 | 21.1987 | 17.4 | 21.67 | 21 | < 100 | 20.3 | 19.9 | 21 | 20.1052 | 19.6 | 22 | 23.3 | [5 empty] | 16
| 1 | 4 | 20 | 22.8711 | 21.58 | 20.3068 | 20.5 | 21.07 | 20.7 | < 100 | 20.4 | 21.4 | 20.68 | 20.1557 | 21.02 | 20.5 | 22.3 | [5 empty] | 20.8147
| 1 | 5 | 21.3 | 22.4508 | 19.84 | 20.1974 | 18.2 | 22.17 | 20.6 | < 100 | 20.9 | 22.2 | 20.4 | 19.4796 | 20.14 | 20.43 | 22.2 | [5 empty] | 47
| 1 | 6 | 21.4 | 22.2263 | 20.93 | 20.5335 | 20.2 | 21.29 | 21.5 | < 100 | 19.75 | 20.6 | 20.2 | 19.9394 | 21.1 | 20.98 | 25.2
| Mean |  | 20.9167 | 22.3842 | 21.045 | 20.4546 | 19.05 | 21.7067 | 21.0167 |  | 20.2917 | 21.1833 | 20.88 | 19.7417 | 20.8817 | 20.8433 | 23.6333
| Median |  | 21 | 22.3385 | 20.84 | 20.358 | 19.05 | 21.545 | 21 |  | 20.35 | 21.35 | 20.84 | 19.7095 | 21.06 | 20.575 | 23.8
| Std Dev. |  | 0.58452 | 0.35137 | 0.96413 | 0.39716 | 1.3035 | 0.58504 | 0.34303 |  | 0.41763 | 0.81833 | 0.55194 | 0.36492 | 0.91421 | 0.59922 | 1.2323
| Rel.Std.Dev. |  | 2.795% | 1.57% | 4.581% | 1.942% | 6.842% | 2.695% | 1.632% |  | 2.058% | 3.863% | 2.643% | 1.848% | 4.378% | 2.875% | 5.214%
| PDM3 |  | 0.49% | 7.541% | 1.107% | -1.73% | -8.478% | 4.285% | 0.9705% |  | -2.513% | 1.771% | 0.3139% | -5.155% | 0.3219% | 0.1377% | 13.54%
| Z-Score (Absolute) |  | 0 | 2.14 | 0.19 | 0.67 | 2.72 | 1.15 | 0.15 | 42.45 | 0.91 | 0.39 | 0.05 | 1.71 | 0.05 | 0.11 | 3.96
| Analytical results for Sn in OREAS 466 (Certified Value 17.6 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 17 | 16.01 | 20.26 | 17.5047 | 3 | 17.5 | 16.4 | 17.5 | 17.2 | 17.9 | 18.2 | 18.4 | 17.2 | 19.0565 | 17.1 | 19.1 | 18.7 | 13.5 |  |  | 1
| 1 | 2 | 17 | 16.41 | 19.06 | 17.4466 | 2 | 17.9 | 16.5 | 16.8 | 17.4 | 18 | 18.6 | 18.2 | 17.1 | 17.4795 | 18.3 | 18.6 | 18.3 | 14.1 |  |  | 0
| 1 | 3 | 17 | 16.56 | 22.14 | 17.9031 | 4 | 17.5 | 16.4 | 17.3 | 16.9 | 17.7 | 16.6 | 18.1 | 17.7 | 19.9299 | 17.2 | 19.1 | 18.1 | 13.2 |  |  | 16
| 1 | 4 | 16 | 15.46 | 23.82 | 17.8625 | 7 | 17.6 | 16.1 | 17.7 | 17.6 | 18.4 | 17.8 | 18 | 17 | 20.048 | 17.4 | 19.8 | 16.9 | 13.9 |  |  | 17.6454
| 1 | 5 | 17 | 16.62 | 19.48 | 17.7484 | 2 | 17.6 | 16.8 | 17.7 | 17.7 | 18.4 | 19.1 | 18 | 17 | 19.6508 | 17.3 | 18.9 | 17.1 | 14.1 |  |  | 48
| 1 | 6 | 18 | 16.26 | 17.96 | 18.1037 | 2 | 17.8 | 16.2 | 17.9 | 16.9 | 17.3 | 16.9 | 18.3 | 17.4 | 19.2341 | 17.5 | 18.2 | 19 | 14.1
| Mean |  | 17 | 16.22 | 20.4533 | 17.7615 | 3.3333 | 17.65 | 16.4 | 17.4833 | 17.2833 | 17.95 | 17.8667 | 18.1667 | 17.2333 | 19.2331 | 17.4667 | 18.95 | 18.0167 | 13.8167
| Median |  | 17 | 16.335 | 19.87 | 17.8054 | 2.5 | 17.6 | 16.4 | 17.6 | 17.3 | 17.95 | 18 | 18.15 | 17.15 | 19.4425 | 17.35 | 19 | 18.2 | 14
| Std Dev. |  | 0.63246 | 0.43243 | 2.1597 | 0.25005 | 1.9664 | 0.16432 | 0.24495 | 0.392 | 0.34303 | 0.42308 | 0.97091 | 0.1633 | 0.27325 | 0.9413 | 0.43205 | 0.53944 | 0.84951 | 0.38166
| Rel.Std.Dev. |  | 3.72% | 2.666% | 10.56% | 1.408% | 58.99% | 0.931% | 1.494% | 2.242% | 1.985% | 2.357% | 5.434% | 0.8989% | 1.586% | 4.894% | 2.474% | 2.847% | 4.715% | 2.762%
| PDM3 |  | -3.658% | -8.078% | 15.91% | 0.6577% | -81.11% | 0.02597% | -7.058% | -0.9186% | -2.052% | 1.726% | 1.254% | 2.954% | -2.335% | 8.998% | -1.013% | 7.393% | 2.104% | -21.7%
| Z-Score (Absolute) |  | 0.75 | 1.79 | 3.83 | 0.26 | 18.88 | 0.11 | 1.55 | 0.11 | 0.38 | 0.51 | 0.4 | 0.8 | 0.44 | 2.21 | 0.13 | 1.84 | 0.6 | 4.97
| Analytical results for Sr in OREAS 466 (Certified Value 107 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 105.12 | 118 | 106.6091 | 105.09 | 107.9587 | 121 | 105 | 109 | 109.5 | 104.5 | 109 | 113 | 102 | 107 | 101.3308 | 120 | 111.63 | 105 | 93 |  | 1
| 1 | 2 | 106.27 | 115 | 108.345 | 104.82 | 109.6648 | 114 | 103 | 105.5 | 105.5 | 101.5 | 107 | 109 | 101 | 106 | 92.6244 | 128 | 106.74 | 105 | 86 |  | 0
| 1 | 3 | 105.94 | 118 | 106.3604 | 104.58 | 108.7501 | 116 | 101 | 107 | 108.5 | 100.5 | 108.5 | 111 | 101 | 106 | 92.6533 | 138 | 109.11 | 106 | 88 |  | 16
| 1 | 4 | 107.46 | 114 | 106.1878 | 107.72 | 108.9337 | 116 | 102 | 105.5 | 115.5 | 104 | 106 | 109 | 100 | 105 | 88.1689 | 121 | 112.19 | 105 | 93 |  | 107.0192
| 1 | 5 | 105.46 | 117 | 108.531 | 101.01 | 109.7951 | 115 | 102 | 108.5 | 110 | 103.5 | 110 | 110 | 100 | 106 | 100.436 | 123 | 109.61 | 106 | 92 |  | 49
| 1 | 6 | 106.21 | 117 | 107.4457 | 98.79 | 109.6813 | 113 | 104 | 104.5 | 112 | 102.5 | 106 | 109 | 102 | 109 | 102.1994 | 129 | 108.85 | 106 | 96
| Mean |  | 106.0767 | 116.5 | 107.2465 | 103.6683 | 109.1306 | 115.8333 | 102.8333 | 106.6667 | 110.1667 | 102.75 | 107.75 | 110.1667 | 101 | 106.5 | 96.2355 | 126.5 | 109.6883 | 105.5 | 91.3333
| Median |  | 106.075 | 117 | 107.0274 | 104.7 | 109.2993 | 115.5 | 102.5 | 106.25 | 109.75 | 103 | 107.75 | 109.5 | 101 | 106 | 96.5447 | 125.5 | 109.36 | 105.5 | 92.5
| Std Dev. |  | 0.8103 | 1.6432 | 1.0207 | 3.2081 | 0.7193 | 2.7869 | 1.472 | 1.8074 | 3.3714 | 1.5411 | 1.6658 | 1.6021 | 0.89443 | 1.3784 | 5.8329 | 6.7157 | 1.9882 | 0.54772 | 3.6697
| Rel.Std.Dev. |  | 0.7639% | 1.41% | 0.9517% | 3.095% | 0.6591% | 2.406% | 1.431% | 1.694% | 3.06% | 1.5% | 1.546% | 1.454% | 0.8856% | 1.294% | 6.061% | 5.309% | 1.813% | 0.5192% | 4.018%
| PDM3 |  | -0.8807% | 8.859% | 0.2124% | -3.131% | 1.973% | 8.236% | -3.911% | -0.3294% | 2.941% | -3.989% | 0.6829% | 2.941% | -5.624% | -0.4851% | -10.08% | 18.2% | 2.494% | -1.42% | -14.66%
| Z-Score (Absolute) |  | 0.11 | 1.89 | 0.11 | 0.58 | 0.47 | 1.77 | 0.74 | 0 | 0.67 | 0.75 | 0.21 | 0.67 | 1.09 | 0.03 | 2.01 | 3.82 | 0.58 | 0.22 | 2.95
| Analytical results for Ta in OREAS 466 (Certified Value 2.56 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.61 | 2.21 | < 5 | 5.1 | < 0.1 | 2.5 | 2.59 | 2.74 | 2.44 | 2.81 | 3 | 1.67 | 2.27 | 2.8805 | 2.71 | 2.98 | 2.23 |  |  |  | 1
| 1 | 2 | 2.77 | 2.23 | < 5 | 3.95 | < 0.1 | 2.4 | 2.53 | 2.55 | 2.54 | 2.67 | 2.3 | 1.651 | 2.25 | 2.9411 | 2.66 | 2.69 | 2.55 |  |  |  | 0
| 1 | 3 | 2.66 | 2.44 | < 5 | 2.9 | 0.1 | 2.4 | 2.49 | 2.53 | 2.43 | 2.81 | 2.4 | 1.66 | 2.44 | 2.9925 | 2.7 | 2.46 | 2.8 |  |  |  | 16
| 1 | 4 | 2.68 | 2.11 | < 5 | 1.55 | 0.2 | 2.4 | 2.49 | 2.67 | 2.47 | 2.75 | 3.1 | 1.651 | 2.33 | 3.0974 | 2.66 | 2.41 | 3.14 |  |  |  | 2.5637
| 1 | 5 | 2.68 | 2.45 | < 5 | 2.38 | < 0.1 | 2.4 | 2.51 | 2.68 | 2.46 | 2.81 | 2.9 | 1.66 | 2.32 | 2.9148 | 2.61 | 2.48 | 3.11 |  |  |  | 50
| 1 | 6 | 2.77 | 2.14 | < 5 | 2.14 | < 0.1 | 2.5 | 2.45 | 2.73 | 2.37 | 2.64 | 2.1 | 1.651 | 2.33 | 2.8956 | 2.61 | 2.67 | 2.45
| Mean |  | 2.695 | 2.2633 |  | 3.0033 | 0.15 | 2.4333 | 2.51 | 2.65 | 2.4517 | 2.7483 | 2.6333 | 1.6572 | 2.3233 | 2.9537 | 2.6583 | 2.615 | 2.7133
| Median |  | 2.68 | 2.22 |  | 2.64 | 0.15 | 2.4 | 2.5 | 2.675 | 2.45 | 2.78 | 2.65 | 1.6555 | 2.325 | 2.928 | 2.66 | 2.575 | 2.675
| Std Dev. |  | 0.063482 | 0.14747 |  | 1.3076 | 0.070711 | 0.05164 | 0.047329 | 0.089666 | 0.055648 | 0.076529 | 0.41793 | 0.007679 | 0.066232 | 0.080696 | 0.042622 | 0.21267 | 0.36784
| Rel.Std.Dev. |  | 2.356% | 6.516% |  | 43.54% | 47.14% | 2.122% | 1.886% | 3.384% | 2.27% | 2.785% | 15.87% | 0.4634% | 2.851% | 2.732% | 1.603% | 8.133% | 13.56%
| PDM3 |  | 5.123% | -11.71% |  | 17.15% | -94.15% | -5.083% | -2.093% | 3.368% | -4.368% | 7.204% | 2.718% | -35.36% | -9.374% | 15.21% | 3.693% | 2.003% | 5.839%
| Z-Score (Absolute) |  | 0.4 | 1.78 | 0.58 | 1.96 | 12.8 | 0.92 | 0.53 | 0.18 | 0.83 | 0.67 | 0.09 | 4.84 | 1.48 | 1.71 | 0.22 | 0 | 0.5
| Analytical results for Tb in OREAS 466 (Certified Value 1.26 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.25 | 1.7415 | 1.979 | 1.2643 | 1.4 | 1.22 | 1.2 | < 500 | 1.15 | 1.4 | 1.27 | 1.3498 | 1.46 | 1.24 | 1.11 | [5 empty] | 1
| 1 | 2 | 1.24 | 1.7209 | 1.97 | 1.2531 | 1.3 | 1.17 | 1.2 | < 500 | 1.15 | 1.4 | 1.26 | 1.3368 | 1.54 | 1.25 | 1.09 | [5 empty] | 0
| 1 | 3 | 1.26 | 1.7112 | 1.945 | 1.2856 | 1.2 | 1.21 | 1.2 | < 500 | 1.11 | 1.3 | 1.27 | 1.3854 | 1.36 | 1.26 | 1.08 | [5 empty] | 16
| 1 | 4 | 1.26 | 1.7346 | 1.702 | 1.2632 | 1.3 | 1.17 | 1.2 | < 500 | 1.14 | 1.4 | 1.26 | 1.4099 | 1.45 | 1.21 | 1.03 | [5 empty] | 1.2646
| 1 | 5 | 1.26 | 1.7462 | 1.872 | 1.2508 | 1.4 | 1.24 | 1.2 | < 500 | 1.17 | 1.5 | 1.268 | 1.3233 | 1.41 | 1.21 | 1.02 | [5 empty] | 51
| 1 | 6 | 1.26 | 1.7287 | 1.912 | 1.2703 | 1.3 | 1.24 | 1.3 | < 500 | 1.15 | 1.3 | 1.27 | 1.3386 | 1.46 | 1.23 | 1.16
| Mean |  | 1.255 | 1.7305 | 1.8967 | 1.2646 | 1.3167 | 1.2083 | 1.2167 |  | 1.145 | 1.3833 | 1.2663 | 1.3573 | 1.4467 | 1.2333 | 1.0817
| Median |  | 1.26 | 1.7317 | 1.9285 | 1.2637 | 1.3 | 1.215 | 1.2 |  | 1.15 | 1.4 | 1.269 | 1.3442 | 1.455 | 1.235 | 1.085
| Std Dev. |  | 0.0083666 | 0.013045 | 0.10318 | 0.012639 | 0.075277 | 0.031885 | 0.040825 |  | 0.019748 | 0.075277 | 0.0049666 | 0.033263 | 0.059889 | 0.020656 | 0.051929
| Rel.Std.Dev. |  | 0.6667% | 0.7538% | 5.44% | 0.9995% | 5.717% | 2.639% | 3.355% |  | 1.725% | 5.442% | 0.3922% | 2.451% | 4.14% | 1.675% | 4.801%
| PDM3 |  | -0.7569% | 36.85% | 49.98% | -0.001041% | 4.12% | -4.447% | -3.788% |  | -9.456% | 9.391% | 0.1393% | 7.333% | 14.4% | -2.47% | -14.46%
| Z-Score (Absolute) |  | 0.08 | 3.44 | 4.67 | 0.01 | 0.37 | 0.43 | 0.37 | 1843.72 | 0.9 | 0.87 | 0 | 0.67 | 1.34 | 0.24 | 1.37
| Analytical results for Te in OREAS 466 (Certified Value 0.3 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.3 | 0.11 | < 0.5 | 0.36305 | < 0.1 | 0.3 | 0.25 | 0.34 | 0.33 | 0.34 | 0.3 | 0.4 | 0.3 | 0.27835 | 0.3 | 0.34 | 0.43 |  |  |  | 1
| 1 | 2 | 0.3 | 0.46 | < 0.5 | 0.34117 | < 0.1 | 0.35 | 0.28 | 0.28 | 0.27 | 0.25 | 0.24 | 0.5 | 0.3 | 0.30433 | 0.3 | 0.35 | 0.46 |  |  |  | 5
| 1 | 3 | 0.3 | 0.15 | < 0.5 | 0.36078 | < 0.1 | 0.32 | 0.28 | 0.36 | 0.21 | 0.29 | 0.32 | 0.4 | 0.3 | 0.31541 | 0.3 | 0.35 | 0.85 |  |  |  | 16
| 1 | 4 | 0.3 | 0.5 | < 0.5 | 0.35444 | < 0.1 | 0.39 | 0.26 | 0.32 | 0.33 | 0.26 | 0.33 | 0.5 | 0.3 | 0.30483 | 0.3 | 0.33 | 0.59 |  |  |  | 0.30302
| 1 | 5 | 0.3 | 0.23 | < 0.5 | 0.35918 | < 0.1 | 0.29 | 0.29 | 0.31 | 0.3 | 0.25 | 0.18 | 0.5 | 0.3 | 0.29384 | 0.3 | 0.31 | 0.83 |  |  |  | 52
| 1 | 6 | 0.3 | 0.35 | < 0.5 | 0.32541 | < 0.1 | 0.3 | 0.27 | 0.36 | 0.34 | 0.22 | 0.17 | 0.4 | 0.3 | 0.31295 | 0.3 | 0.31 | 0.75
| Mean |  | 0.3 | 0.3 |  | 0.35067 |  | 0.325 | 0.27167 | 0.32833 | 0.29667 | 0.26833 | 0.25667 | 0.45 | 0.3 | 0.30162 | 0.3 | 0.33167 | 0.65167
| Median |  | 0.3 | 0.29 |  | 0.35681 |  | 0.31 | 0.275 | 0.33 | 0.315 | 0.255 | 0.27 | 0.45 | 0.3 | 0.30458 | 0.3 | 0.335 | 0.67
| Std Dev. |  | 0 | 0.16223 |  | 0.014633 |  | 0.038341 | 0.01472 | 0.031252 | 0.049666 | 0.041673 | 0.070616 | 0.054772 | 0 | 0.013698 | 0 | 0.018348 | 0.18465
| Rel.Std.Dev. |  | 0% | 54.08% |  | 4.173% |  | 11.8% | 5.418% | 9.518% | 16.74% | 15.53% | 27.51% | 12.17% | 0% | 4.542% | 0% | 5.532% | 28.34%
| PDM3 |  | -0.9953% | -0.9953% |  | 15.73% |  | 7.255% | -10.35% | 8.355% | -2.095% | -11.45% | -15.3% | 48.51% | -0.9953% | -0.4612% | -0.9953% | 9.455% | 115.1%
| Z-Score (Absolute) |  | 0 | 0 | 1.12 | 1.14 | 5.62 | 0.56 | 0.64 | 0.64 | 0.07 | 0.71 | 0.97 | NA | NA | 0.04 | NA | 0.71 | 7.9
Results from laboratories 16, 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 466 (Certified Value 32.4 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 32.54 | 32.35 | 33.6124 | 34.74 | 30.537 | 34.6 | 29.91 | 32.4 | 31.9 | 32.6 | 33.3 | 30.7 | 32.6 | 34 | 30.1 | 29.547 | 36.23 | 31.02 | 40.5 | 35 | 1
| 1 | 2 | 31.63 | 33.52 | 33.2138 | 32.07 | 30.8407 | 33.2 | 27.81 | 32.2 | 33.2 | 30.5 | 32.2 | 31.3 | 32.2 | 35 | 30.7 | 29.4687 | 37.24 | 31.07 | 40.2 | 38 | 0
| 1 | 3 | 33.61 | 32.27 | 33.5806 | 32.28 | 31.7375 | 32 | 29.21 | 31.9 | 31.7 | 31.5 | 31.4 | 31.7 | 28.4 | 34.1 | 30.3 | 29.8857 | 35.89 | 31.88 | 40 | 34.5 | 16
| 1 | 4 | 32.24 | 30.31 | 33.4795 | 36.06 | 30.9301 | 32.4 | 28.11 | 31.9 | 33.2 | 33 | 31.9 | 30.9 | 31.4 | 35.3 | 31.1 | 30.7416 | 36.21 | 31.14 | 38.8 | 36.2 | 32.4156
| 1 | 5 | 32.82 | 32.34 | 33.2659 | 33.04 | 30.3603 | 34.1 | 28.71 | 31.2 | 33.3 | 33.8 | 32.6 | 32.1 | 33 | 34.4 | 29.7 | 30.2524 | 36.21 | 31.34 | 37.4 | 36.8 | 53
| 1 | 6 | 34.09 | 31.55 | 33.9232 | 33.86 | 30.8495 | 33.2 | 29.01 | 32.6 | 33.7 | 33.7 | 31.3 | 30.8 | 29.3 | 35.5 | 31.1 | 30.3737 | 35.48 | 31.99 | 42.4 | 36.3
| Mean |  | 32.8217 | 32.0567 | 33.5126 | 33.675 | 30.8758 | 33.25 | 28.7933 | 32.0333 | 32.8333 | 32.5167 | 32.1167 | 31.25 | 31.15 | 34.7167 | 30.5 | 30.0449 | 36.21 | 31.4067 | 39.8833 | 36.1333
| Median |  | 32.68 | 32.305 | 33.53 | 33.45 | 30.8451 | 33.2 | 28.86 | 32.05 | 33.2 | 32.8 | 32.05 | 31.1 | 31.8 | 34.7 | 30.5 | 30.0691 | 36.21 | 31.24 | 40.1 | 36.25
| Std Dev. |  | 0.90198 | 1.064 | 0.25863 | 1.5351 | 0.47495 | 0.98336 | 0.76267 | 0.49261 | 0.82381 | 1.2952 | 0.75741 | 0.55767 | 1.8802 | 0.63692 | 0.56569 | 0.49835 | 0.58251 | 0.42491 | 1.6833 | 1.258
| Rel.Std.Dev. |  | 2.748% | 3.319% | 0.7717% | 4.558% | 1.538% | 2.957% | 2.649% | 1.538% | 2.509% | 3.983% | 2.358% | 1.785% | 6.036% | 1.835% | 1.855% | 1.659% | 1.609% | 1.353% | 4.221% | 3.482%
| PDM3 |  | 1.253% | -1.107% | 3.384% | 3.885% | -4.75% | 2.574% | -11.17% | -1.179% | 1.289% | 0.3118% | -0.9222% | -3.596% | -3.904% | 7.099% | -5.91% | -7.314% | 11.71% | -3.113% | 23.04% | 11.47%
| Z-Score (Absolute) |  | 0.29 | 0.15 | 0.68 | 0.78 | 0.82 | 0.53 | 2.01 | 0.16 | 0.29 | 0.11 | 0.11 | 0.61 | 0.67 | 1.37 | 1.04 | 1.3 | 2.22 | 0.52 | 4.32 | 2.18
| Analytical results for Ti in OREAS 466 (Certified Value 0.448 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.464 | 0.4087 | 0.50619 | 0.4481 | 0.45862 | 0.15 | 0.458 | 0.432 | 0.506 | 0.478 | 0.476 | 0.461 | 0.375 | 0.397 | 0.57337 | 0.342 | 0.4803 | 0.47 | 0.48 |  | 1
| 1 | 2 | 0.4705 | 0.413 | 0.50174 | 0.41589 | 0.43268 | 0.0773 | 0.448 | 0.424 | 0.487 | 0.498 | 0.451 | 0.5 | 0.3733 | 0.405 | 0.59493 | 0.377 | 0.4495 | 0.47 | 0.45 |  | 0
| 1 | 3 | 0.4566 | 0.3973 | 0.50809 | 0.45079 | 0.46657 | 0.293 | 0.436 | 0.418 | 0.475 | 0.483 | 0.474 | 0.44 | 0.3755 | 0.418 | 0.56761 | 0.389 | 0.4773 | 0.48 | 0.48 |  | 16
| 1 | 4 | 0.4562 | 0.4032 | 0.50901 | 0.42731 | 0.49409 | 0.365 | 0.456 | 0.433 | 0.537 | 0.495 | 0.464 | 0.426 | 0.375 | 0.387 | 0.59627 | 0.367 | 0.4866 | 0.47 | 0.5 |  | 0.44801
| 1 | 5 | 0.4537 | 0.4026 | 0.50416 | 0.42148 | 0.45702 | 0.189 | 0.449 | 0.414 | 0.495 | 0.494 | 0.468 | 0.458 | 0.37407 | 0.404 | 0.55656 | 0.332 | 0.4534 | 0.48 | 0.48 |  | 54
| 1 | 6 | 0.4749 | 0.4229 | 0.50433 | 0.43152 | 0.46238 | 0.0688 | 0.454 | 0.439 | 0.499 | 0.465 | 0.47 | 0.433 | 0.3754 | 0.4 | 0.57106 | 0.37 | 0.4463 | 0.47 | 0.52
| Mean |  | 0.46265 | 0.40795 | 0.50559 | 0.43252 | 0.46189 | 0.19052 | 0.45017 | 0.42667 | 0.49983 | 0.4855 | 0.46717 | 0.453 | 0.37471 | 0.40183 | 0.57663 | 0.36283 | 0.46557 | 0.47333 | 0.485
| Median |  | 0.4603 | 0.40595 | 0.50526 | 0.42942 | 0.4605 | 0.1695 | 0.4515 | 0.428 | 0.497 | 0.4885 | 0.469 | 0.449 | 0.375 | 0.402 | 0.57222 | 0.3685 | 0.46535 | 0.47 | 0.48
| Std Dev. |  | 0.0086234 | 0.0091014 | 0.0027106 | 0.014163 | 0.01973 | 0.11851 | 0.0079603 | 0.0096264 | 0.021094 | 0.012661 | 0.0089981 | 0.026818 | 0.00085567 | 0.010226 | 0.01579 | 0.021628 | 0.017745 | 0.005164 | 0.023452
| Rel.Std.Dev. |  | 1.864% | 2.231% | 0.5361% | 3.275% | 4.271% | 62.2% | 1.768% | 2.256% | 4.22% | 2.608% | 1.926% | 5.92% | 0.2284% | 2.545% | 2.738% | 5.961% | 3.812% | 1.091% | 4.835%
| PDM3 |  | 3.267% | -8.942% | 12.85% | -3.459% | 3.099% | -57.48% | 0.4809% | -4.765% | 11.57% | 8.368% | 4.275% | 1.113% | -16.36% | -10.31% | 28.71% | -19.01% | 3.918% | 5.652% | 8.256%
| Z-Score (Absolute) |  | 0.02 | 1.24 | 1 | 0.67 | 0 | 6.23 | 0.27 | 0.81 | 0.87 | 0.54 | 0.12 | 0.2 | 2 | 1.38 | 2.63 | 2.27 | 0.08 | 0.26 | 0.53
| Analytical results for Tl in OREAS 466 (Certified Value 0.79 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 0.77 | 0.773 | 0.87761 | 0.664 | 0.8582 | 0.8 | 0.78 | 0.7 | 0.78 | 0.73 | 0.69 | 0.84 | 0.8 | 0.78 | 0.78662 | 0.86 | 0.79 | 0.77 | 0.78 |  | 1
| 1 | 2 | 0.78 | 0.867 | 0.8672 | 0.633 | 0.86274 | 0.77 | 0.79 | 0.75 | 0.72 | 0.77 | 0.65 | 0.78 | 0.8 | 0.77 | 0.78588 | 0.92 | 0.8 | 0.74 | 0.9 |  | 6
| 1 | 3 | 0.79 | 0.823 | 0.86906 | 0.626 | 0.84958 | 0.76 | 0.8 | 0.68 | 0.76 | 0.68 | 0.73 | 0.67 | 0.8 | 0.77 | 0.79047 | 0.83 | 0.83 | 0.74 | 0.82 |  | 16
| 1 | 4 | 0.8 | 0.77 | 0.87414 | 0.677 | 0.91893 | 0.76 | 0.8 | 0.72 | 0.77 | 0.72 | 0.7 | 0.77 | 0.8 | 0.78 | 0.78403 | 0.89 | 0.79 | 0.73 | 0.85 |  | 0.78802
| 1 | 5 | 0.79 | 0.787 | 0.8868 | 0.621 | 0.93174 | 0.76 | 0.78 | 0.71 | 0.78 | 0.73 | 0.74 | 0.82 | 0.8 | 0.76 | 0.79154 | 0.88 | 0.81 | 0.71 | 0.84 |  | 55
| 1 | 6 | 0.8 | 0.796 | 0.87793 | 0.655 | 0.88025 | 0.76 | 0.82 | 0.75 | 0.8 | 0.7 | 0.68 | 0.72 | 0.8 | 0.78 | 0.74907 | 0.85 | 0.81 | 0.79 | 0.8
| Mean |  | 0.78833 | 0.80267 | 0.87546 | 0.646 | 0.88357 | 0.76833 | 0.795 | 0.71833 | 0.76833 | 0.72167 | 0.69833 | 0.76667 | 0.8 | 0.77333 | 0.78127 | 0.87167 | 0.805 | 0.74667 | 0.83167
| Median |  | 0.79 | 0.7915 | 0.87587 | 0.644 | 0.8715 | 0.76 | 0.795 | 0.715 | 0.775 | 0.725 | 0.695 | 0.775 | 0.8 | 0.775 | 0.78625 | 0.87 | 0.805 | 0.74 | 0.83
| Std Dev. |  | 0.01169 | 0.036849 | 0.0070739 | 0.022627 | 0.034102 | 0.016021 | 0.015166 | 0.027869 | 0.027142 | 0.030605 | 0.033116 | 0.06314 | 1.2162e-16 | 0.008165 | 0.016028 | 0.031885 | 0.015166 | 0.028752 | 0.042151
| Rel.Std.Dev. |  | 1.483% | 4.591% | 0.808% | 3.503% | 3.86% | 2.085% | 1.908% | 3.88% | 3.533% | 4.241% | 4.742% | 8.236% | 1.52e-14% | 1.056% | 2.052% | 3.658% | 1.884% | 3.851% | 5.068%
| PDM3 |  | 0.04006% | 1.859% | 11.1% | -18.02% | 12.13% | -2.498% | 0.8861% | -8.843% | -2.498% | -8.42% | -11.38% | -2.709% | 1.521% | -1.863% | -0.8565% | 10.62% | 2.155% | -5.247% | 5.539%
| Z-Score (Absolute) |  | 0.26 | 0.59 | 2.27 | 3.04 | 2.46 | 0.21 | 0.41 | 1.36 | 0.21 | 1.29 | 1.83 | 0.25 | NA | 0.09 | 0.09 | 2.18 | 0.64 | 0.71 | 1.26
Results from laboratory 16 were removed due to their 0.001 ppm reading resolution.
| Analytical results for Tm in OREAS 466 (Certified Value 0.27 ppm) | [21 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 0.28 | 0.28083 | 0.281 | 0.29364 | 0.4 | 0.22 | 0.3 | 0.26 | 0.3 | 0.272 | 0.46136 | 0.24 | 0.3 | 0.24 | [6 empty] | 1
| 1 | 2 | 0.29 | 0.2775 | 0.362 | 0.2863 | 0.3 | 0.17 | 0.3 | 0.27 | 0.3 | 0.262 | 0.47148 | 0.24 | 0.28 | 0.22 | [6 empty] | 0
| 1 | 3 | 0.28 | 0.31566 | 0.315 | 0.29282 | 0.3 | 0.22 | 0.3 | 0.27 | 0.3 | 0.272 | 0.48028 | 0.25 | 0.31 | 0.22 | [6 empty] | 16
| 1 | 4 | 0.29 | 0.27972 | 0.282 | 0.28126 | 0.4 | 0.18 | 0.3 | 0.27 | 0.3 | 0.268 | 0.49373 | 0.24 | 0.29 | 0.21 | [6 empty] | 0.27431
| 1 | 5 | 0.31 | 0.2921 | 0.277 | 0.28171 | 0.4 | 0.18 | 0.3 | 0.27 | 0.3 | 0.276 | 0.48575 | 0.24 | 0.29 | 0.21 | [6 empty] | 56
| 1 | 6 | 0.31 | 0.2879 | 0.398 | 0.29069 | 0.4 | 0.19 | 0.3 | 0.26 | 0.3 | 0.272 | 0.47368 | 0.25 | 0.28 | 0.24
| Mean |  | 0.29333 | 0.28895 | 0.31917 | 0.28773 | 0.36667 | 0.19333 | 0.3 | 0.26667 | 0.3 | 0.27033 | 0.47771 | 0.24333 | 0.29167 | 0.22333
| Median |  | 0.29 | 0.28436 | 0.2985 | 0.28849 | 0.4 | 0.185 | 0.3 | 0.27 | 0.3 | 0.272 | 0.47698 | 0.24 | 0.29 | 0.22
| Std Dev. |  | 0.013663 | 0.014189 | 0.050364 | 0.0054722 | 0.05164 | 0.021602 | 0 | 0.005164 | 0 | 0.0048028 | 0.011404 | 0.005164 | 0.01169 | 0.013663
| Rel.Std.Dev. |  | 4.658% | 4.91% | 15.78% | 1.902% | 14.08% | 11.17% | 0% | 1.936% | 0% | 1.777% | 2.387% | 2.122% | 4.008% | 6.118%
| PDM3 |  | 6.937% | 5.339% | 16.35% | 4.896% | 33.67% | -29.52% | 9.367% | -2.785% | 9.367% | -1.448% | 74.15% | -11.29% | 6.329% | -18.58%
| Z-Score (Absolute) |  | 0.18 | 0.04 | 1.01 | 0 | NA | 3.02 | NA | 0.67 | NA | 0.56 | 6.08 | 1.42 | 0.13 | 2.06
Results from laboratories 7, 9 and 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 466 (Certified Value 3.04 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 3.1 | 3.02 | 3.2438 | 3.702 | 2.8553 | 3.3 | 3.42 | 3.1 | 2.7 | 3 | 2.9 | 2.9 | 3.3 | 2.9 | 2.71 | 3.0665 | 3.59 | 3 | 2.95 | 2.84 | 1
| 1 | 2 | 3.07 | 3.02 | 3.2053 | 3.57 | 2.8169 | 3.2 | 3.49 | 3.1 | 2.8 | 2.8 | 2.9 | 2.9 | 3.3 | 2.8 | 2.74 | 3.145 | 3.52 | 2.97 | 2.91 | 3.07 | 0
| 1 | 3 | 3.08 | 2.96 | 3.2329 | 3.815 | 2.9961 | 3.2 | 3.52 | 3.1 | 2.7 | 2.9 | 2.8 | 2.9 | 3 | 2.9 | 2.79 | 3.2104 | 3.71 | 3.17 | 2.82 | 2.81 | 16
| 1 | 4 | 3.2 | 2.81 | 3.2309 | 2.941 | 2.8901 | 3.2 | 3.53 | 3 | 2.8 | 3.2 | 2.9 | 2.9 | 3.3 | 2.8 | 2.74 | 3.2193 | 3.76 | 3 | 2.79 | 2.98 | 3.0374
| 1 | 5 | 3.1 | 2.95 | 3.2989 | 3.066 | 2.7881 | 3.4 | 3.49 | 3.1 | 2.8 | 3.1 | 2.9 | 2.9 | 3.4 | 2.9 | 2.73 | 3.1255 | 3.66 | 2.98 | 2.73 | 2.93 | 57
| 1 | 6 | 3.17 | 2.94 | 3.2659 | 3.421 | 2.9392 | 3.8 | 3.47 | 3.1 | 2.9 | 3.1 | 2.8 | 2.8 | 3.1 | 2.9 | 2.74 | 3.1231 | 3.9 | 2.96 | 2.99 | 2.92
| Mean |  | 3.12 | 2.95 | 3.2463 | 3.4192 | 2.881 | 3.35 | 3.4867 | 3.0833 | 2.7833 | 3.0167 | 2.8667 | 2.8833 | 3.2333 | 2.8667 | 2.7417 | 3.1483 | 3.69 | 3.0133 | 2.865 | 2.925
| Median |  | 3.1 | 2.955 | 3.2383 | 3.4955 | 2.8727 | 3.25 | 3.49 | 3.1 | 2.8 | 3.05 | 2.9 | 2.9 | 3.3 | 2.9 | 2.74 | 3.1353 | 3.685 | 2.99 | 2.865 | 2.925
| Std Dev. |  | 0.052536 | 0.076942 | 0.032415 | 0.35009 | 0.077622 | 0.23452 | 0.039328 | 0.040825 | 0.075277 | 0.1472 | 0.05164 | 0.040825 | 0.15055 | 0.05164 | 0.026394 | 0.057892 | 0.13357 | 0.078401 | 0.10075 | 0.094393
| Rel.Std.Dev. |  | 1.684% | 2.608% | 0.9985% | 10.24% | 2.694% | 7.001% | 1.128% | 1.324% | 2.705% | 4.879% | 1.801% | 1.416% | 4.656% | 1.801% | 0.9627% | 1.839% | 3.62% | 2.602% | 3.516% | 3.227%
| PDM3 |  | 2.72% | -2.877% | 6.877% | 12.57% | -5.151% | 10.29% | 14.79% | 1.512% | -8.365% | -0.6825% | -5.621% | -5.072% | 6.451% | -5.621% | -9.736% | 3.651% | 21.49% | -0.7922% | -5.676% | -3.7%
| Z-Score (Absolute) |  | 0.48 | 0.3 | 1.05 | 1.84 | 0.61 | 1.52 | 2.14 | 0.31 | 1.05 | 0.01 | 0.67 | 0.6 | 0.99 | 0.67 | 1.24 | 0.61 | 3.07 | 0.01 | 0.68 | 0.41
| Analytical results for V in OREAS 466 (Certified Value 128 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 129 | 127 | 123.2126 | 122.4 | 127.4158 | 60 | 127 | 122 | 132 | 124 | 128 | 134 | 124 | 134 | 137.9959 | 134 | 122 | 132 | 131 |  | 1
| 1 | 2 | 125 | 128 | 125.3306 | 122.74 | 125.7573 | 49 | 126 | 122 | 126 | 124 | 127 | 138 | 127.72 | 133 | 140.9569 | 145 | 127 | 130 | 131 |  | 0
| 1 | 3 | 128 | 130 | 126.7345 | 122.92 | 128.274 | 86 | 126 | 121 | 129 | 122 | 127 | 129 | 127.72 | 133 | 139.7781 | 152 | 127 | 130 | 128 |  | 16
| 1 | 4 | 130 | 128 | 123.7121 | 123.74 | 126.9332 | 92 | 127 | 121 | 139 | 127 | 127 | 128 | 121.54 | 133 | 145.5833 | 153 | 127 | 121 | 133 |  | 127.5664
| 1 | 5 | 131 | 130 | 123.8084 | 120.52 | 125.6819 | 77 | 126 | 121 | 131 | 124 | 128 | 133 | 124.63 | 133 | 136.2054 | 154 | 132 | 131 | 129 |  | 58
| 1 | 6 | 135 | 128 | 122.9599 | 120.61 | 129.4781 | 46 | 126 | 122 | 131 | 126 | 128 | 127 | 121.54 | 132 | 139.7432 | 145 | 129 | 132 | 132
| Mean |  | 129.6667 | 128.5 | 124.293 | 122.155 | 127.2567 | 68.3333 | 126.3333 | 121.5 | 131.3333 | 124.5 | 127.5 | 131.5 | 124.525 | 133 | 140.0438 | 147.1667 | 127.3333 | 129.3333 | 130.6667
| Median |  | 129.5 | 128 | 123.7602 | 122.57 | 127.1745 | 68.5 | 126 | 121.5 | 131 | 124 | 127.5 | 131 | 124.315 | 133 | 139.7606 | 148.5 | 127 | 130.5 | 131
| Std Dev. |  | 3.3267 | 1.2247 | 1.4529 | 1.3084 | 1.4716 | 19.4388 | 0.5164 | 0.54772 | 4.3205 | 1.7607 | 0.54772 | 4.2308 | 2.7757 | 0.63246 | 3.1814 | 7.5741 | 3.266 | 4.1793 | 1.8619
| Rel.Std.Dev. |  | 2.566% | 0.9531% | 1.169% | 1.071% | 1.156% | 28.45% | 0.4088% | 0.4508% | 3.29% | 1.414% | 0.4296% | 3.217% | 2.229% | 0.4755% | 2.272% | 5.147% | 2.565% | 3.231% | 1.425%
| PDM3 |  | 1.646% | 0.7318% | -2.566% | -4.242% | -0.2428% | -46.43% | -0.9666% | -4.756% | 2.953% | -2.404% | -0.05208% | 3.084% | -2.384% | 4.259% | 9.781% | 15.36% | -0.1827% | 1.385% | 2.43%
| Z-Score (Absolute) |  | 0.46 | 0.21 | 0.68 | 1.14 | 0.05 | 12.6 | 0.25 | 1.28 | 0.82 | 0.64 | 0 | 0.85 | 0.63 | 1.17 | 2.67 | 4.19 | 0.04 | 0.39 | 0.67
| Analytical results for W in OREAS 466 (Certified Value 90 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 0
1
| 1 | 1 | 91.3 | 85.78 | 68.6552 | 19.5077 | 88.9531 | 0.7 | 93.8 | 92.2 | 98.1 | 95.7 | 98 | 84.7 | 80 | 85.9 | 89.9308 | 93.4 | 120 | 74.6 |  |  | 1
| 1 | 2 | 94.9 | 87.87 | 64.1394 | 18.9555 | 90.7525 | 0.6 | 94.5 | 96.9 | 92.5 | 101.5 | 92.3 | 90 | 80 | 87.6 | 95.9258 | 92.3 | 120 | 81.9 |  |  | 0
| 1 | 3 | 94.4 | 84.46 | 68.0234 | 19.5812 | 90.9588 | 2.2 | 91.6 | 89.4 | 90.7 | 97.2 | 100.5 | 77 | 80 | 89 | 95.3218 | 93 | 126 | 73.9 |  |  | 16
| 1 | 4 | 93.8 | 79.41 | 65.433 | 18.6633 | 91.63 | 1.9 | 93.3 | 94 | 99 | 93.2 | 96.7 | 85.2 | 80 | 87.6 | 98.5175 | 95.3 | 113 | 83.3 |  |  | 90.1577
| 1 | 5 | 91.7 | 84.19 | 66.479 | 18.8405 | 88.2391 | 0.7 | 91.9 | 95.6 | 98.1 | 99.1 | 100 | 89.5 | 80 | 86.3 | 95.9007 | 88.6 | 107 | 75.4 |  |  | 59
| 1 | 6 | 95.9 | 83.89 | 65.9645 | 19.3757 | 93.2561 | 1.1 | 93.2 | 95.3 | 98.6 | 95 | 93 | 80.1 | 80 | 87.8 | 92.8629 | 94 | 118 | 76.1
| Mean |  | 93.6667 | 84.2667 | 66.4491 | 19.154 | 90.6316 | 1.2 | 93.05 | 93.9 | 96.1667 | 96.95 | 96.75 | 84.4167 | 80 | 87.3667 | 94.7433 | 92.7667 | 117.3333 | 77.5333
| Median |  | 94.1 | 84.325 | 66.2218 | 19.1656 | 90.8556 | 0.9 | 93.25 | 94.65 | 98.1 | 96.45 | 97.35 | 84.95 | 80 | 87.6 | 95.6113 | 93.2 | 119 | 75.75
| Std Dev. |  | 1.8184 | 2.7945 | 1.6702 | 0.38344 | 1.8192 | 0.68702 | 1.1113 | 2.7203 | 3.5987 | 2.9938 | 3.4657 | 5.1254 | 0 | 1.1183 | 2.9669 | 2.2792 | 6.5625 | 4.0183
| Rel.Std.Dev. |  | 1.941% | 3.316% | 2.514% | 2.002% | 2.007% | 57.25% | 1.194% | 2.897% | 3.742% | 3.088% | 3.582% | 6.072% | 0% | 1.28% | 3.132% | 2.457% | 5.593% | 5.183%
| PDM3 |  | 3.892% | -6.534% | -26.3% | -78.76% | 0.5256% | -98.67% | 3.208% | 4.151% | 6.665% | 7.534% | 7.312% | -6.368% | -11.27% | -3.096% | 5.086% | 2.894% | 30.14% | -14%
| Z-Score (Absolute) |  | 0.26 | 0.97 | 3.31 | 9.5 | 0.14 | 11.85 | 0.18 | 0.29 | 0.58 | 0.69 | 0.66 | 0.95 | 1.53 | 0.57 | 0.4 | 0.14 | 3.36 | 1.85
| Analytical results for Y in OREAS 466 (Certified Value 20.3 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 20.12 | 18.92 | 21.4635 | 20.5376 | 20.2516 | 26.7 | 20.28 | 18.8 | 19.3 | 21.7 | 21.1 | 19.7 | 19.9 | 19.0099 | 20.8 | 33.7941 | 23.4 | 19.92 | 21.2 | 19.1 | 1
| 1 | 2 | 20.2 | 18.63 | 21.209 | 20.6153 | 19.8214 | 26.4 | 20.82 | 19.1 | 19.2 | 21.2 | 21.5 | 20.2 | 20 | 18.0198 | 21.3 | 32.3393 | 22.7 | 20.38 | 20.5 | 19 | 0
| 1 | 3 | 20.4 | 20.39 | 21.4845 | 21.3788 | 20.6107 | 25 | 20.28 | 20.1 | 19.3 | 37.5 | 20.3 | 19.8 | 17.7 | 19.703 | 21.1 | 32.0957 | 23.1 | 20.35 | 20.9 | 17.7 | 16
| 1 | 4 | 20.45 | 16.49 | 21.3787 | 20.7188 | 20.1533 | 26.3 | 19.86 | 18 | 19.1 | 24.3 | 20.7 | 19.8 | 19 | 18.6 | 20.9 | 32.7287 | 23 | 19.61 | 18.9 | 18 | 20.2983
| 1 | 5 | 20.21 | 21.81 | 21.923 | 20.7835 | 20.1947 | 26.2 | 20.36 | 19 | 19.4 | 22.2 | 21.4 | 20.4 | 20.1 | 18.8119 | 21.2 | 34.1446 | 23.4 | 20.05 | 19.3 | 18.2 | 60
| 1 | 6 | 20.3 | 18.57 | 21.7038 | 19.9682 | 20.9916 | 25.5 | 20.04 | 19.1 | 19 | 22.5 | 22.3 | 19.2 | 20.4 | 18.7129 | 21.2 | 33.9251 | 23.5 | 19.95 | 21.7 | 18.5
| Mean |  | 20.28 | 19.135 | 21.5271 | 20.6671 | 20.3372 | 26.0167 | 20.2733 | 19.0167 | 19.2167 | 24.9 | 21.2167 | 19.85 | 19.5167 | 18.8096 | 21.0833 | 33.1712 | 23.1833 | 20.0433 | 20.4167 | 18.4167
| Median |  | 20.255 | 18.775 | 21.474 | 20.6671 | 20.2231 | 26.25 | 20.28 | 19.05 | 19.25 | 22.35 | 21.25 | 19.8 | 19.95 | 18.7624 | 21.15 | 33.2614 | 23.25 | 20 | 20.7 | 18.35
| Std Dev. |  | 0.12696 | 1.8082 | 0.2518 | 0.45401 | 0.40759 | 0.63692 | 0.32611 | 0.67355 | 0.1472 | 6.2626 | 0.69402 | 0.41833 | 1.0068 | 0.55051 | 0.19408 | 0.88865 | 0.30605 | 0.28953 | 1.0998 | 0.55648
| Rel.Std.Dev. |  | 0.6261% | 9.449% | 1.17% | 2.197% | 2.004% | 2.448% | 1.609% | 3.542% | 0.766% | 25.15% | 3.271% | 2.107% | 5.159% | 2.927% | 0.9205% | 2.679% | 1.32% | 1.445% | 5.387% | 3.022%
| PDM3 |  | -0.09012% | -5.731% | 6.054% | 1.817% | 0.1918% | 28.17% | -0.123% | -6.314% | -5.329% | 22.67% | 4.524% | -2.209% | -3.851% | -7.334% | 3.868% | 63.42% | 14.21% | -1.256% | 0.5832% | -9.27%
| Z-Score (Absolute) |  | 0.02 | 0.79 | 0.82 | 0.24 | 0.02 | 3.85 | 0.02 | 0.87 | 0.74 | 3.1 | 0.61 | 0.31 | 0.53 | 1.01 | 0.52 | 8.67 | 1.94 | 0.18 | 0.07 | 1.28
| Analytical results for Yb in OREAS 466 (Certified Value 1.83 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.94 | 2.183 | 2.405 | 1.7786 | 2.4 | 1.7 | 1.7 | 1.77 | 1.9 | 1.892 | 3.0624 | 1.7 | 1.91 | 1.7 | [6 empty] | 1
| 1 | 2 | 1.94 | 2.1583 | 2.611 | 1.8064 | 2.5 | 1.48 | 1.8 | 1.78 | 1.9 | 1.908 | 3.0635 | 1.8 | 1.99 | 1.7 | [6 empty] | 0
| 1 | 3 | 1.91 | 2.1201 | 2.868 | 1.8179 | 2.3 | 1.56 | 1.7 | 1.75 | 1.6 | 1.908 | 3.2131 | 1.6 | 2 | 1.7 | [6 empty] | 16
| 1 | 4 | 2.47 | 2.1148 | 2.016 | 1.8748 | 2.3 | 1.51 | 1.7 | 1.76 | 1.8 | 1.806 | 3.1897 | 1.7 | 2.09 | 1.7 | [6 empty] | 1.8252
| 1 | 5 | 1.95 | 2.1736 | 2.255 | 1.8528 | 2.5 | 1.51 | 1.7 | 1.84 | 1.9 | 1.84 | 3.1059 | 1.7 | 1.92 | 1.7 | [6 empty] | 61
| 1 | 6 | 1.94 | 2.1519 | 2.616 | 1.8806 | 2.3 | 1.67 | 1.8 | 1.78 | 1.8 | 1.81 | 3.037 | 1.7 | 1.95 | 1.8
| Mean |  | 2.025 | 2.1503 | 2.4618 | 1.8352 | 2.3833 | 1.5717 | 1.7333 | 1.78 | 1.8167 | 1.8607 | 3.1119 | 1.7 | 1.9767 | 1.7167
| Median |  | 1.94 | 2.1551 | 2.508 | 1.8354 | 2.35 | 1.535 | 1.7 | 1.775 | 1.85 | 1.866 | 3.0847 | 1.7 | 1.97 | 1.7
| Std Dev. |  | 0.21843 | 0.027755 | 0.30192 | 0.04068 | 0.098319 | 0.09196 | 0.05164 | 0.031623 | 0.1169 | 0.047844 | 0.073133 | 0.063246 | 0.066232 | 0.040825
| Rel.Std.Dev. |  | 10.79% | 1.291% | 12.26% | 2.217% | 4.125% | 5.851% | 2.979% | 1.777% | 6.435% | 2.571% | 2.35% | 3.72% | 3.351% | 2.378%
| PDM3 |  | 10.95% | 17.81% | 34.88% | 0.548% | 30.58% | -13.89% | -5.033% | -2.476% | -0.4672% | 1.943% | 70.5% | -6.859% | 8.299% | -5.946%
| Z-Score (Absolute) |  | 0.86 | 1.46 | 2.97 | 0.06 | 2.59 | 1.33 | 0.55 | 0.33 | 0.15 | 0.06 | 6.11 | 0.71 | 0.62 | 0.63
| Analytical results for Zn in OREAS 466 (Certified Value 41.8 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 40 | 35 | 58.2708 | 57.4033 | 42.828 | 46 | 40.2 | 40 | 41 | 40 | 41 | 44.3 | 44 | 40 | 41.5537 | 46 | 43 | 41 | 52 |  | 1
| 1 | 2 | 43 | 35 | 54.5714 | 50.7507 | 43.3292 | 46 | 39.5 | 39 | 39 | 40 | 40 | 45.9 | 43 | 40 | 43.6745 | 41 | 40 | 40 | 67 |  | 0
| 1 | 3 | 41 | 34 | 54.9968 | 53.8333 | 42.9026 | 46 | 39.8 | 39 | 40 | 39 | 43 | 43.2 | 44 | 42 | 41.7578 | 43 | 43 | 40 | 53 |  | 16
| 1 | 4 | 42 | 34 | 53.2434 | 70.5727 | 43.0011 | 46 | 40.4 | 39 | 42 | 41 | 40 | 45.6 | 44 | 40 | 43.4159 | 40 | 43 | 39 | 52 |  | 41.8345
| 1 | 5 | 41 | 35 | 54.1888 | 64.0107 | 43.2369 | 46 | 39.4 | 39 | 41 | 40 | 41 | 44.3 | 46 | 40 | 41.0344 | 41 | 45 | 41 | 53 |  | 62
| 1 | 6 | 41 | 40 | 53.9488 | 49.8553 | 43.821 | 46 | 39.9 | 39 | 41 | 40 | 41 | 45 | 45 | 40 | 43.0462 | 39 | 41 | 41 | 55
| Mean |  | 41.3333 | 35.5 | 54.87 | 57.7377 | 43.1865 | 46 | 39.8667 | 39.1667 | 40.6667 | 40 | 41 | 44.7167 | 44.3333 | 40.3333 | 42.4137 | 41.6667 | 42.5 | 40.3333 | 55.3333
| Median |  | 41 | 35 | 54.3801 | 55.6183 | 43.119 | 46 | 39.85 | 39 | 41 | 40 | 41 | 44.65 | 44 | 40 | 42.402 | 41 | 43 | 40.5 | 53
| Std Dev. |  | 1.0328 | 2.2583 | 1.768 | 8.1289 | 0.36578 | 0 | 0.38816 | 0.40825 | 1.0328 | 0.63246 | 1.0954 | 0.99079 | 1.0328 | 0.8165 | 1.1015 | 2.5033 | 1.7607 | 0.8165 | 5.8195
| Rel.Std.Dev. |  | 2.499% | 6.361% | 3.222% | 14.08% | 0.847% | 0% | 0.9736% | 1.042% | 2.54% | 1.581% | 2.672% | 2.216% | 2.33% | 2.024% | 2.597% | 6.008% | 4.143% | 2.024% | 10.52%
| PDM3 |  | -1.198% | -15.14% | 31.16% | 38.01% | 3.232% | 9.957% | -4.704% | -6.377% | -2.791% | -4.385% | -1.995% | 6.89% | 5.973% | -3.588% | 1.385% | -0.4011% | 1.591% | -3.588% | 32.27%
| Z-Score (Absolute) |  | 0.13 | 2.49 | 5.34 | 6.5 | 0.61 | 1.75 | 0.73 | 1.01 | 0.4 | 0.67 | 0.27 | 1.23 | 1.08 | 0.54 | 0.3 | 0 | 0.34 | 0.54 | 5.53
| Analytical results for Zr in OREAS 466 (Certified Value 161 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 160.8 | 150.2 | 172.2536 | 165.3058 | 154.9444 | 137 | 153.8 | 168.5 | 168 | 164 | 166.5 | 169 | 163.6 | 156 | 259.1319 | 185 | 165.6 | 183 | 159 |  | 1
| 1 | 2 | 161.9 | 147.3 | 172.7172 | 158.6632 | 152.5078 | 114 | 151.7 | 164 | 163.5 | 162.5 | 166.5 | 170.4 | 162.8 | 159 | 279.5298 | 182 | 160.6 | 183 | 160 |  | 0
| 1 | 3 | 162 | 147.7 | 178.5165 | 161.448 | 155.8905 | 156 | 153.3 | 166.5 | 163.5 | 158.5 | 166 | 162.3 | 157.6 | 158 | 277.9201 | 194 | 166.3 | 183 | 149 |  | 16
| 1 | 4 | 165.2 | 143.2 | 178.5491 | 154.1182 | 157.9756 | 168 | 153.5 | 165 | 176.5 | 163 | 164 | 159.8 | 162 | 160 | 295.4907 | 189 | 171.3 | 183 | 160 |  | 161.3336
| 1 | 5 | 159.6 | 144.7 | 178.8333 | 154.7572 | 154.0291 | 147 | 154.6 | 168 | 171.5 | 161 | 172.5 | 169.6 | 157 | 154 | 272.934 | 188 | 160.2 | 185 | 153 |  | 63
| 1 | 6 | 163.9 | 151.7 | 174.8842 | 205.8727 | 154.6694 | 97 | 156.1 | 164 | 173.5 | 160.5 | 162.5 | 156.8 | 157.2 | 159 | 259.0022 | 185 | 161.8 | 185 | 159
| Mean |  | 162.2333 | 147.4667 | 175.959 | 166.6942 | 155.0028 | 136.5 | 153.8333 | 166 | 169.4167 | 161.5833 | 166.3333 | 164.65 | 160.0333 | 157.6667 | 274.0014 | 187.1667 | 164.3 | 183.6667 | 156.6667
| Median |  | 161.95 | 147.5 | 176.7003 | 160.0556 | 154.8069 | 142 | 153.65 | 165.75 | 169.75 | 161.75 | 166.25 | 165.65 | 159.8 | 158.5 | 275.4271 | 186.5 | 163.7 | 183 | 159
| Std Dev. |  | 2.0363 | 3.2042 | 3.0629 | 19.6446 | 1.8392 | 26.6439 | 1.461 | 1.9748 | 5.3518 | 1.9854 | 3.4157 | 5.782 | 3.0787 | 2.2509 | 13.8197 | 4.1673 | 4.2764 | 1.0328 | 4.5898
| Rel.Std.Dev. |  | 1.255% | 2.173% | 1.741% | 11.78% | 1.187% | 19.52% | 0.9498% | 1.19% | 3.159% | 1.229% | 2.053% | 3.512% | 1.924% | 1.428% | 5.044% | 2.227% | 2.603% | 0.5623% | 2.93%
| PDM3 |  | 0.5577% | -8.595% | 9.065% | 3.323% | -3.924% | -15.39% | -4.649% | 2.892% | 5.01% | 0.1548% | 3.099% | 2.056% | -0.8059% | -2.273% | 69.84% | 16.01% | 1.839% | 13.84% | -2.893%
| Z-Score (Absolute) |  | 0.21 | 1.71 | 1.19 | 0.24 | 0.95 | 2.83 | 1.06 | 0.17 | 0.52 | 0.28 | 0.21 | 0.04 | 0.43 | 0.67 | 11.15 | 2.32 | 0 | 1.97 | 0.78
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 466 (Indicative Value < 1 ppm) | [19 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | [16 empty] | 1
| No | No | 2.23 | 2.25 | [16 empty] | ppm
|  |  | PF*OES | PF*MS | [16 empty] | 2
2
| 1 | 1 | < 5 | < 1 | [16 empty] | 1
| 1 | 2 | < 5 | < 1 | [16 empty] | 4
| 1 | 3 | < 5 | < 1 | [16 empty] | 16
| 1 | 4 | < 5 | < 1 | [16 empty] | < 1
| 1 | 5 | < 5 | < 1 | [16 empty] | 10
| 1 | 6 | < 5 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Al in OREAS 466 (Certified Value 8.34 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 2
3
| 1 | 1 | 8.4119 | 8.1184 | 8.6607 | 8.6787 | 8.45 | 8.24 | 8.39 | 8.14 | 8.39 | 7.7616 | 8.2616 | [7 empty] | 1
| 1 | 2 | 8.4477 | 8.1611 | 8.6361 | 8.6074 | 8.44 | 8.54 | 8.09 | 8.14 | 8.36 | 7.7834 | 8.4193 | [7 empty] | 0
| 1 | 3 | 8.4719 | 8.1441 | 8.6166 | 8.569 | 8.35 | 8.23 | 8.05 | 8.15 | 8.35 | 7.6977 | 8.2902 | [7 empty] | 16
| 1 | 4 | 8.2926 | 8.1742 | 8.9309 | 8.6336 | 8.45 | 8.19 | 8.48 | 8.14 | 8.17 | 7.7211 | 8.3355 | [7 empty] | 8.345
| 1 | 5 | 8.4782 | 8.193 | 8.7721 | 8.5968 | 8.38 | 8.38 | 8.29 | 8.15 | 8.39 | 7.6066 | 8.3005 | [7 empty] | 65
| 1 | 6 | 8.4977 | 8.1483 | 8.6895 | 8.5948 | 8.52 | 8.79 | 8.36 | 8.13 | 8.37 | 7.7491 | 8.3012
| Mean |  | 8.4333 | 8.1565 | 8.7176 | 8.6134 | 8.4317 | 8.395 | 8.2767 | 8.1417 | 8.3383 | 7.7199 | 8.318
| Median |  | 8.4598 | 8.1547 | 8.6751 | 8.6021 | 8.445 | 8.31 | 8.325 | 8.14 | 8.365 | 7.7351 | 8.3008
| Std Dev. |  | 0.074987 | 0.02585 | 0.11778 | 0.038218 | 0.059805 | 0.23262 | 0.17178 | 0.0075277 | 0.084004 | 0.063174 | 0.05496
| Rel.Std.Dev. |  | 0.8892% | 0.3169% | 1.351% | 0.4437% | 0.7093% | 2.771% | 2.075% | 0.09246% | 1.007% | 0.8183% | 0.6607%
| PDM3 |  | 1.059% | -2.258% | 4.466% | 3.217% | 1.039% | 0.5997% | -0.8183% | -2.436% | -0.07936% | -7.49% | -0.3226%
| Z-Score (Absolute) |  | 0.67 | 1.29 | 2.69 | 1.95 | 0.66 | 0.4 | 0.44 | 1.4 | 0 | 4.39 | 0.14
| Analytical results for Al2O3 in OREAS 466 (Certified Value 15.77 wt.%) | [19 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 2
3
| 1 | 1 | 15.8943 | 15.3397 | 16.3644 | 16.3983 | 15.966 | 15.569 | 15.853 | 15.381 | 15.853 | 14.6655 | 15.6103 | [7 empty] | 1
| 1 | 2 | 15.9619 | 15.4204 | 16.3179 | 16.2637 | 15.947 | 16.136 | 15.286 | 15.381 | 15.796 | 14.7067 | 15.9082 | [7 empty] | 0
| 1 | 3 | 16.0077 | 15.3882 | 16.281 | 16.191 | 15.777 | 15.551 | 15.21 | 15.399 | 15.777 | 14.5448 | 15.6643 | [7 empty] | 16
| 1 | 4 | 15.6689 | 15.4451 | 16.875 | 16.3133 | 15.966 | 15.475 | 16.023 | 15.381 | 15.437 | 14.589 | 15.75 | [7 empty] | 15.7678
| 1 | 5 | 16.0196 | 15.4807 | 16.5749 | 16.2437 | 15.834 | 15.834 | 15.664 | 15.399 | 15.853 | 14.3727 | 15.6838 | [7 empty] | 66
| 1 | 6 | 16.0564 | 15.3963 | 16.4188 | 16.2398 | 16.099 | 16.609 | 15.796 | 15.362 | 15.815 | 14.6419 | 15.685
| Mean |  | 15.9348 | 15.4117 | 16.472 | 16.275 | 15.9315 | 15.8623 | 15.6387 | 15.3838 | 15.7552 | 14.5868 | 15.7169
| Median |  | 15.9848 | 15.4083 | 16.3916 | 16.2537 | 15.9565 | 15.7015 | 15.73 | 15.381 | 15.8055 | 14.6155 | 15.6844
| Std Dev. |  | 0.14169 | 0.048844 | 0.22255 | 0.072213 | 0.11318 | 0.43959 | 0.32471 | 0.013862 | 0.15881 | 0.11937 | 0.10385
| Rel.Std.Dev. |  | 0.8892% | 0.3169% | 1.351% | 0.4437% | 0.7104% | 2.771% | 2.076% | 0.09011% | 1.008% | 0.8183% | 0.6607%
| PDM3 |  | 1.059% | -2.258% | 4.466% | 3.217% | 1.038% | 0.5996% | -0.8189% | -2.435% | -0.08008% | -7.49% | -0.3226%
| Z-Score (Absolute) |  | 0.67 | 1.29 | 2.69 | 1.95 | 0.66 | 0.4 | 0.44 | 1.4 | 0 | 4.39 | 0.14
| Analytical results for As in OREAS 466 (Certified Value 472 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | [12 empty] | 0
0
| 1 | 1 | 473 | 500 | 500 | 490 | 460 | 503 | [12 empty] | 1
| 1 | 2 | 473 | 500 | 400 | 490 | 420 | 503 | [12 empty] | 0
| 1 | 3 | 447 | 500 | 400 | 480 | 445 | 498 | [12 empty] | 16
| 1 | 4 | 468 | 500 | 500 | 500 | 405 | 498 | [12 empty] | 472.25
| 1 | 5 | 459 | 500 | 500 | 480 | 415 | 499 | [12 empty] | 67
| 1 | 6 | 453 | 500 | 400 | 490 | 460 | 492
| Mean |  | 462.1667 | 500 | 450 | 488.3333 | 434.1667 | 498.8333
| Median |  | 463.5 | 500 | 450 | 490 | 432.5 | 498.5
| Std Dev. |  | 10.8888 | 0 | 54.7723 | 7.5277 | 23.9618 | 4.0702
| Rel.Std.Dev. |  | 2.356% | 0% | 12.17% | 1.542% | 5.519% | 0.8159%
| PDM3 |  | -2.135% | 5.876% | -4.711% | 3.406% | -8.064% | 5.629%
| Z-Score (Absolute) |  | 0.37 | 0.69 | 0.7 | 0.37 | 1.15 | 0.66
| Analytical results for B in OREAS 466 (Certified Value 0.138 wt.%) | [19 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.07 | 2.17 | 2.23 | 2.25 | [13 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | [13 empty] | 3
3
| 1 | 1 | 0.1471 | 0.121 | 0.139 | 0.147 | 0.14 | [13 empty] | 1
| 1 | 2 | 0.1469 | 0.115 | 0.135 | 0.154 | 0.14 | [13 empty] | 0
| 1 | 3 | 0.1477 | 0.122 | 0.138 | 0.155 | 0.14 | [13 empty] | 16
| 1 | 4 | 0.1434 | 0.123 | 0.137 | 0.144 | 0.13 | [13 empty] | 0.1383
| 1 | 5 | 0.1485 | 0.116 | 0.138 | 0.147 | 0.13 | [13 empty] | 68
| 1 | 6 | 0.1484 | 0.125 | 0.136 | 0.155 | 0.14
| Mean |  | 0.147 | 0.12033 | 0.13717 | 0.15033 | 0.13667
| Median |  | 0.1474 | 0.1215 | 0.1375 | 0.1505 | 0.14
| Std Dev. |  | 0.0018804 | 0.0039833 | 0.001472 | 0.0048854 | 0.005164
| Rel.Std.Dev. |  | 1.279% | 3.31% | 1.073% | 3.25% | 3.779%
| PDM3 |  | 6.291% | -12.99% | -0.8195% | 8.701% | -1.181%
| Z-Score (Absolute) |  | 0.67 | 1.15 | 0 | 0.9 | 0.03
| Analytical results for Ba in OREAS 466 (Certified Value 719 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 0
0
| 1 | 1 | 736 | 729.824 | 718.6374 | 713.1466 | 714 | 714 | 700 | 712 | 710 | 674 | 715.085 | [7 empty] | 1
| 1 | 2 | 732 | 738.96 | 769.3006 | 693.1179 | 684 | 720 | 700 | 728 | 735 | 672 | 730.1 | [7 empty] | 0
| 1 | 3 | 740 | 732.56 | 714.0951 | 733.0651 | 728 | 702 | 700 | 713 | 735 | 667 | 721.475 | [7 empty] | 16
| 1 | 4 | 803 | 733.456 | 730.9489 | 695.6139 | 696 | 701 | 700 | 716 | 690 | 665 | 720.705 | [7 empty] | 718.8531
| 1 | 5 | 745 | 734.856 | 721.6154 | 711.6997 | 684 | 710 | 700 | 721 | 705 | 671 | 726.455 | [7 empty] | 69
| 1 | 6 | 739 | 731.216 | 729.7529 | 696.8448 | 690 | 695 | 700 | 732 | 735 | 671 | 721.43
| Mean |  | 749.1667 | 733.4787 | 730.7251 | 707.248 | 699.3333 | 707 | 700 | 720.3333 | 718.3333 | 670 | 722.5417
| Median |  | 739.5 | 733.008 | 725.6842 | 704.2723 | 693 | 706 | 700 | 718.5 | 722.5 | 671 | 721.4525
| Std Dev. |  | 26.7239 | 3.202 | 19.9736 | 15.2602 | 17.9183 | 9.2952 | 0 | 8.2138 | 19.4079 | 3.3466 | 5.1728
| Rel.Std.Dev. |  | 3.567% | 0.4365% | 2.733% | 2.158% | 2.562% | 1.315% | 0% | 1.14% | 2.702% | 0.4995% | 0.7159%
| PDM3 |  | 4.217% | 2.035% | 1.652% | -1.614% | -2.715% | -1.649% | -2.623% | 0.2059% | -0.07231% | -6.796% | 0.5131%
| Z-Score (Absolute) |  | 1.65 | 0.78 | 0.63 | 0.67 | 1.1 | 0.68 | NA | 0.06 | 0.06 | 2.72 | 0.18
Results from laboratory 15 were removed due to their 100 ppm reading resolution.
| Analytical results for BaO in OREAS 466 (Certified Value 803 ppm) | [19 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 0
0
| 1 | 1 | 821.7 | 814.8485 | 802.3587 | 796.2282 | 797.2 | 797.2 | 780 | 794.9 | 792.7 | 752.52 | 798.3924 | [7 empty] | 1
| 1 | 2 | 817.3 | 825.0488 | 858.9241 | 773.8661 | 763.7 | 803.9 | 780 | 812.8 | 820.6 | 750.29 | 815.1566 | [7 empty] | 0
| 1 | 3 | 826.2 | 817.9032 | 797.2872 | 818.4671 | 812.8 | 783.8 | 780 | 796.1 | 820.6 | 744.71 | 805.5268 | [7 empty] | 16
| 1 | 4 | 896.5 | 818.9036 | 816.1045 | 776.653 | 777.1 | 782.7 | 780 | 799.4 | 770.4 | 742.47 | 804.6671 | [7 empty] | 802.5995
| 1 | 5 | 831.8 | 820.4667 | 805.6836 | 794.6128 | 763.7 | 792.7 | 780 | 805 | 787.1 | 749.17 | 811.087 | [7 empty] | 70
| 1 | 6 | 825.1 | 816.4027 | 814.7691 | 778.0272 | 770.4 | 776 | 780 | 817.3 | 820.6 | 749.17 | 805.4766
| Mean |  | 836.4333 | 818.9289 | 815.8545 | 789.6424 | 780.8167 | 789.3833 | 780 | 804.25 | 802 | 748.055 | 806.7178
| Median |  | 825.65 | 818.4034 | 810.2264 | 786.32 | 773.75 | 788.25 | 780 | 802.2 | 806.65 | 749.17 | 805.5017
| Std Dev. |  | 29.8188 | 3.575 | 22.3005 | 17.038 | 19.9981 | 10.371 | 0 | 9.1797 | 21.6561 | 3.7363 | 5.7755
| Rel.Std.Dev. |  | 3.565% | 0.4365% | 2.733% | 2.158% | 2.561% | 1.314% | 0% | 1.141% | 2.7% | 0.4995% | 0.7159%
| PDM3 |  | 4.216% | 2.035% | 1.652% | -1.614% | -2.714% | -1.647% | -2.816% | 0.2056% | -0.0747% | -6.796% | 0.5131%
| Z-Score (Absolute) |  | 1.65 | 0.78 | 0.63 | 0.67 | 1.1 | 0.68 | NA | 0.06 | 0.06 | 2.72 | 0.18
Results from laboratory 15 were removed due to their 100 ppm reading resolution.
| Analytical results for Be in OREAS 466 (Certified Value 3.12 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.09 | 2.17 | 2.23 | 2.25 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 3 | 6.3555 | 2.7677 | 3.1463 | 3 | 3 | 3 | 3.2 | < 5 | [9 empty] | 1
| 1 | 2 | 4 | 9.242 | 3.2367 | 3.1608 | 3 | 3 | 3 | 3.4 | < 5 | [9 empty] | 0
| 1 | 3 | 3 | 8.9074 | 3.0599 | 3.3392 | 4 | 4 | 3 | 3.3 | < 5 | [9 empty] | 16
| 1 | 4 | 3 | 8.6755 | 2.6641 | 3.1699 | < 3 | 3 | 3 | 3 | < 5 | [9 empty] | 3.1236
| 1 | 5 | 3 | 4.6628 | 2.6881 | 2.9956 | < 3 | 3 | 3 | 3.2 | < 5 | [9 empty] | 71
| 1 | 6 | 3 | 5.9379 | 2.6155 | 3.0277 | 4 | 2 | 3 | 3.9 | < 5
| Mean |  | 3.1667 | 7.2969 | 2.8387 | 3.1399 | 3.5 | 3 | 3 | 3.3333
| Median |  | 3 | 7.5155 | 2.7279 | 3.1535 | 3.5 | 3 | 3 | 3.25
| Std Dev. |  | 0.40825 | 1.8947 | 0.25115 | 0.12207 | 0.57735 | 0.63246 | 0 | 0.30768
| Rel.Std.Dev. |  | 12.89% | 25.97% | 8.847% | 3.888% | 16.5% | 21.08% | 0% | 9.23%
| PDM3 |  | 1.378% | 133.6% | -9.122% | 0.5221% | 12.05% | -3.957% | -3.957% | 6.714%
| Z-Score (Absolute) |  | 0.67 | 17.38 | 0.65 | 0.57 | 0.67 | 0 | 0 | 1.35 | 2.02
| Analytical results for Bi in OREAS 466 (Certified Value 22 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.17 | 2.23 | 2.25 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | [12 empty] | 1
2
| 1 | 1 | 21 |  | 23 | 22.1 | 21.2 | 22.4 | [12 empty] | 1
| 1 | 2 | 21 |  | 20 | 21.7 | 22.6 | 22.3 | [12 empty] | 0
| 1 | 3 | 21 | 4.4165 | 22 | 21.3 | 23.9 | 24.5 | [12 empty] | 16
| 1 | 4 | 21 |  | 21 | 21.9 | 22.3 | 22.4 | [12 empty] | 21.9567
| 1 | 5 | 21 | 6.7055 | 21 | 22.4 | 23 | 21.1 | [12 empty] | 72
| 1 | 6 | 23 |  | 22 | 21.6 | 20.9 | 24.1
| Mean |  | 21.3333 | 5.561 | 21.5 | 21.8333 | 22.3167 | 22.8
| Median |  | 21 | 5.561 | 21.5 | 21.8 | 22.45 | 22.4
| Std Dev. |  | 0.8165 | 1.6186 | 1.0488 | 0.38816 | 1.1232 | 1.2681
| Rel.Std.Dev. |  | 3.827% | 29.11% | 4.878% | 1.778% | 5.033% | 5.562%
| PDM3 |  | -2.839% | -74.67% | -2.08% | -0.5617% | 1.64% | 3.841%
| Z-Score (Absolute) |  | 0.46 | 22.09 | 0.23 | 0.23 | 0.89 | 1.55
| Analytical results for Ca in OREAS 466 (Certified Value 0.245 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.26 | 0.2336 | 0.23899 | 0.068397 | 0.23 | 0.23 | 0.27 | 0.2 | 0.3 | 0.2751 | 0.22585 | [7 empty] | 1
| 1 | 2 | 0.25 | 0.2402 | 0.23746 | 0.094867 | 0.26 | 0.23 | 0.24 | 0.2 | 0.3 | 0.2743 | 0.23665 | [7 empty] | 0
| 1 | 3 | 0.26 | 0.2429 | 0.23255 | 0.4511 | 0.27 | 0.23 | 0.25 | 0.2 | 0.3 | 0.4754 | 0.23154 | [7 empty] | 16
| 1 | 4 | 0.26 | 0.3392 | 0.2385 | 0.046369 | 0.25 | 0.23 | 0.28 | 0.2 | 0.3 | 0.2903 | 0.23359 | [7 empty] | 0.24507
| 1 | 5 | 0.25 | 0.2352 | 0.23711 | 0.14466 | 0.22 | 0.24 | 0.27 | 0.2 | 0.2 | 0.2222 | 0.22888 | [7 empty] | 73
| 1 | 6 | 0.25 | 0.2378 | 0.23491 |  | 0.25 | 0.24 | 0.27 | 0.2 | 0.3 | 0.2222 | 0.22897
| Mean |  | 0.255 | 0.25482 | 0.23659 | 0.16108 | 0.24667 | 0.23333 | 0.26333 | 0.2 | 0.28333 | 0.29325 | 0.23091
| Median |  | 0.255 | 0.239 | 0.23728 | 0.094867 | 0.25 | 0.23 | 0.27 | 0.2 | 0.3 | 0.2747 | 0.23026
| Std Dev. |  | 0.0054772 | 0.041475 | 0.0024323 | 0.16622 | 0.018619 | 0.005164 | 0.015055 | 3.0405e-17 | 0.040825 | 0.093779 | 0.0038443
| Rel.Std.Dev. |  | 2.148% | 16.28% | 1.028% | 103.2% | 7.548% | 2.213% | 5.717% | 1.52e-14% | 14.41% | 31.98% | 1.665%
| PDM3 |  | 4.05% | 3.975% | -3.463% | -34.27% | 0.6498% | -4.791% | 7.45% | -18.39% | 15.61% | 19.66% | -5.778%
| Z-Score (Absolute) |  | 0.42 | 0.41 | 0.51 | 4.33 | 0 | 0.67 | 0.84 | NA | NA | 2.36 | 0.8
Results from laboratories 17 and 23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for CaO in OREAS 466 (Certified Value 0.343 wt.%) | [19 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.364 | 0.32685 | 0.33439 | 0.0957 | 0.322 | 0.322 | 0.378 | 0.28 | 0.42 | 0.38492 | 0.31601 | [7 empty] | 1
| 1 | 2 | 0.35 | 0.33609 | 0.33225 | 0.13274 | 0.364 | 0.322 | 0.336 | 0.28 | 0.42 | 0.3838 | 0.33112 | [7 empty] | 0
| 1 | 3 | 0.364 | 0.33987 | 0.32538 | 0.63118 | 0.378 | 0.322 | 0.35 | 0.28 | 0.42 | 0.66518 | 0.32397 | [7 empty] | 16
| 1 | 4 | 0.364 | 0.47461 | 0.33371 | 0.06488 | 0.35 | 0.322 | 0.392 | 0.28 | 0.42 | 0.40619 | 0.32683 | [7 empty] | 0.34291
| 1 | 5 | 0.35 | 0.32909 | 0.33176 | 0.20241 | 0.308 | 0.336 | 0.378 | 0.28 | 0.28 | 0.3109 | 0.32025 | [7 empty] | 74
| 1 | 6 | 0.35 | 0.33273 | 0.32869 |  | 0.35 | 0.336 | 0.378 | 0.28 | 0.42 | 0.3109 | 0.32038
| Mean |  | 0.357 | 0.35654 | 0.33103 | 0.22538 | 0.34533 | 0.32667 | 0.36867 | 0.28 | 0.39667 | 0.41032 | 0.32309
| Median |  | 0.357 | 0.33441 | 0.33201 | 0.13274 | 0.35 | 0.322 | 0.378 | 0.28 | 0.42 | 0.38436 | 0.32217
| Std Dev. |  | 0.0076681 | 0.058032 | 0.0034032 | 0.23257 | 0.026067 | 0.0072296 | 0.021078 | 0 | 0.057155 | 0.13122 | 0.0053789
| Rel.Std.Dev. |  | 2.148% | 16.28% | 1.028% | 103.2% | 7.548% | 2.213% | 5.717% | 0% | 14.41% | 31.98% | 1.665%
| PDM3 |  | 4.11% | 3.976% | -3.463% | -34.27% | 0.7074% | -4.736% | 7.512% | -18.35% | 15.68% | 19.66% | -5.778%
| Z-Score (Absolute) |  | 0.42 | 0.41 | 0.51 | 4.33 | 0 | 0.67 | 0.84 | NA | NA | 2.36 | 0.8
Results from laboratories 17 and 23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 466 (Certified Value < 10 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.07 | 2.15 | 2.17 | 2.23 | 2.25 | [12 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 5 | < 2 | < 50 | < 10 | < 0.5 | < 10 | [12 empty] | 1
| 1 | 2 | < 5 | < 2 | < 50 | < 10 | < 0.5 | < 10 | [12 empty] | 1
| 1 | 3 | < 5 | < 2 | < 50 | < 10 | < 0.5 | < 10 | [12 empty] | 16
| 1 | 4 | < 5 | < 2 | < 50 | < 10 | < 0.5 | < 10 | [12 empty] | < 10
| 1 | 5 | < 5 | < 2 | < 50 | < 10 | < 0.5 | < 10 | [12 empty] | 75
| 1 | 6 | < 5 | < 2 | < 50 | < 10 | < 0.5 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.42 | 0.93 | 7.16 | 0.42 | 1.18 | 0.42
| Analytical results for Ce in OREAS 466 (Certified Value 0.036 wt.%) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | wt.%
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | 3
3
| 1 | 1 | 0.03628 | 0.0352 | 0.040085 | 0.03756 | 0.037196 | 0.0347 | 0.040667 | 0.03347 | 0.0347 | 0.0324 | 0.0334 | 0.0373 | 0.0349 | 0.0366 | 0.0319 | 0.04095 | 0.03351 | 0.034592 | 1
| 1 | 2 | 0.03519 | 0.0356 | 0.040306 | 0.0364 | 0.034217 | 0.0334 | 0.039902 | 0.0334 | 0.0323 | 0.0334 | 0.0366 | 0.0371 | 0.0358 | 0.0372 | 0.032 | 0.04104 | 0.03619 | 0.037248 | 0
| 1 | 3 | 0.03572 | 0.0359 | 0.040433 | 0.03383 | 0.037796 | 0.0338 | 0.039834 | 0.03314 | 0.0332 | 0.0328 | 0.0333 | 0.0371 | 0.0352 | 0.0369 | 0.0323 | 0.0408 | 0.03375 | 0.035107 | 16
| 1 | 4 | 0.03503 | 0.038 | 0.040534 | 0.0411 | 0.035993 | 0.0344 | 0.040376 | 0.03243 | 0.0351 | 0.0321 | 0.0349 | 0.0368 | 0.0357 | 0.0369 | 0.0306 | 0.04044 | 0.03548 | 0.035418 | 0.035574
| 1 | 5 | 0.03492 | 0.0354 | 0.040176 | 0.03591 | 0.036456 | 0.0342 | 0.041138 | 0.03296 | 0.0344 | 0.0328 | 0.0353 | 0.0378 | 0.035 | 0.0371 | 0.0311 | 0.04059 | 0.03594 | 0.03515 | 76
| 1 | 6 | 0.03553 | 0.0361 | 0.040424 | 0.03612 | 0.035456 | 0.0336 | 0.039666 | 0.03226 | 0.0339 | 0.0336 | 0.0337 | 0.037 | 0.0355 | 0.0363 | 0.0319 | 0.0406 | 0.03537 | 0.034672
| Mean |  | 0.035445 | 0.036033 | 0.040326 | 0.03682 | 0.036186 | 0.034017 | 0.040264 | 0.032943 | 0.033933 | 0.03285 | 0.034533 | 0.037183 | 0.03535 | 0.036833 | 0.031633 | 0.040737 | 0.03504 | 0.035364
| Median |  | 0.03536 | 0.03575 | 0.040365 | 0.03626 | 0.036224 | 0.034 | 0.040139 | 0.03305 | 0.03415 | 0.0328 | 0.0343 | 0.0371 | 0.03535 | 0.0369 | 0.0319 | 0.0407 | 0.035425 | 0.035129
| Std Dev. |  | 0.00050852 | 0.0010172 | 0.00017028 | 0.0024208 | 0.0012753 | 0.00049967 | 0.00056787 | 0.00050107 | 0.0010367 | 0.00057184 | 0.0013033 | 0.00034303 | 0.00037283 | 0.00033267 | 0.00064395 | 0.00023226 | 0.0011349 | 0.0009738
| Rel.Std.Dev. |  | 1.435% | 2.823% | 0.4223% | 6.575% | 3.524% | 1.469% | 1.41% | 1.521% | 3.055% | 1.741% | 3.774% | 0.9225% | 1.055% | 0.9032% | 2.036% | 0.5702% | 3.239% | 2.754%
| PDM3 |  | -0.3622% | 1.292% | 13.36% | 3.503% | 1.72% | -4.377% | 13.18% | -7.394% | -4.612% | -7.657% | -2.925% | 4.524% | -0.6292% | 3.541% | -11.08% | 14.51% | -1.501% | -0.5886%
| Z-Score (Absolute) |  | 0.02 | 0.3 | 2.33 | 0.67 | 0.37 | 0.66 | 2.3 | 1.17 | 0.7 | 1.21 | 0.41 | 0.84 | 0.03 | 0.68 | 1.79 | 2.53 | 0.17 | 0.02
| Analytical results for CeO2 in OREAS 466 (Certified Value 0.044 wt.%) | [19 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | wt.%
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | 3
3
| 1 | 1 | 0.044566 | 0.04324 | 0.04924 | 0.046139 | 0.045692 | 0.042625 | 0.049955 | 0.041115 | 0.04263 | 0.0398 | 0.04103 | 0.04582 | 0.04287 | 0.044959 | 0.039186 | 0.050303 | 0.041164 | 0.042492 | 1
| 1 | 2 | 0.043227 | 0.04373 | 0.049512 | 0.044714 | 0.042033 | 0.041029 | 0.049016 | 0.041029 | 0.03968 | 0.04103 | 0.04496 | 0.04557 | 0.04398 | 0.045696 | 0.039309 | 0.050414 | 0.044456 | 0.045756 | 0
| 1 | 3 | 0.043878 | 0.0441 | 0.049668 | 0.041557 | 0.046429 | 0.04152 | 0.048932 | 0.040709 | 0.04078 | 0.04029 | 0.04091 | 0.04557 | 0.04324 | 0.045328 | 0.039677 | 0.050119 | 0.041459 | 0.043125 | 16
| 1 | 4 | 0.043031 | 0.04668 | 0.049792 | 0.050487 | 0.044213 | 0.042257 | 0.049598 | 0.039837 | 0.04312 | 0.03943 | 0.04287 | 0.04521 | 0.04385 | 0.045328 | 0.037589 | 0.049676 | 0.043584 | 0.043507 | 0.043699
| 1 | 5 | 0.042896 | 0.04349 | 0.049353 | 0.044112 | 0.044783 | 0.042011 | 0.050534 | 0.040488 | 0.04226 | 0.04029 | 0.04336 | 0.04643 | 0.04299 | 0.045574 | 0.038203 | 0.049861 | 0.044149 | 0.043179 | 77
| 1 | 6 | 0.043645 | 0.04435 | 0.049656 | 0.04437 | 0.043554 | 0.041274 | 0.048726 | 0.039628 | 0.04164 | 0.04127 | 0.0414 | 0.04545 | 0.04361 | 0.044591 | 0.039186 | 0.049873 | 0.043449 | 0.042591
| Mean |  | 0.043541 | 0.044265 | 0.049537 | 0.04523 | 0.04445 | 0.041786 | 0.04946 | 0.040468 | 0.041685 | 0.040352 | 0.042422 | 0.045675 | 0.043423 | 0.045246 | 0.038858 | 0.050041 | 0.043043 | 0.043442
| Median |  | 0.043436 | 0.043915 | 0.049584 | 0.044542 | 0.044498 | 0.041765 | 0.049307 | 0.040599 | 0.04195 | 0.04029 | 0.042135 | 0.04557 | 0.043425 | 0.045328 | 0.039186 | 0.049996 | 0.043516 | 0.043152
| Std Dev. |  | 0.00062446 | 0.0012493 | 0.00020918 | 0.0029735 | 0.0015666 | 0.00061357 | 0.00069756 | 0.00061572 | 0.0012747 | 0.00070212 | 0.0015994 | 0.00041951 | 0.0004592 | 0.00040864 | 0.00079105 | 0.00028555 | 0.0013941 | 0.0011962
| Rel.Std.Dev. |  | 1.434% | 2.822% | 0.4223% | 6.574% | 3.524% | 1.468% | 1.41% | 1.522% | 3.058% | 1.74% | 3.77% | 0.9185% | 1.058% | 0.9031% | 2.036% | 0.5706% | 3.239% | 2.754%
| PDM3 |  | -0.3625% | 1.295% | 13.36% | 3.503% | 1.72% | -4.377% | 13.18% | -7.394% | -4.609% | -7.66% | -2.923% | 4.522% | -0.6306% | 3.54% | -11.08% | 14.51% | -1.501% | -0.5886%
| Z-Score (Absolute) |  | 0.02 | 0.3 | 2.33 | 0.67 | 0.37 | 0.66 | 2.3 | 1.17 | 0.7 | 1.21 | 0.41 | 0.84 | 0.03 | 0.68 | 1.79 | 2.53 | 0.17 | 0.02
| Analytical results for Co in OREAS 466 (Certified Value 3.26 ppm) | [19 empty] | Certified
Co | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.05; 2.05: 2.07; 2.06: 2.09; 2.07: 2.15; 2.08: 2.17; 2.09: 2.23; 2.10: 2.25; 2.11: 2.26 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*OES/MS; 2.06: BF*MS; 2.07: PF*OES; 2.08: PF*OES/MS; 2.09: PF*MS; 2.10: BF*MS; 2.11: BF*MS | 2
2
| 1 | 1 | 2.01: 5; 2.02: < 5; 2.04: 6.0927; 2.05: 3.5; 2.06: 3.1; 2.07: < 10; 2.08: 3; 2.09: 4; 2.10: 3.2; 2.11: 3.1286 | 1
| 1 | 2 | 2.01: 5; 2.02: < 5; 2.04: 19.7669; 2.05: 3.4; 2.06: 3; 2.07: < 10; 2.08: 3; 2.09: 4; 2.10: 2.4; 2.11: 3.3933 | 0
| 1 | 3 | 2.01: 5; 2.02: < 5; 2.04: 22.6046; 2.05: 3; 2.06: 3.1; 2.07: < 10; 2.08: 3; 2.09: 4; 2.10: 3.8; 2.11: 3.0907 | 16
| 1 | 4 | 2.01: 6; 2.02: < 5; 2.04: 7.6239; 2.05: 3.3; 2.06: 3.2; 2.07: < 10; 2.08: 3; 2.09: 3; 2.10: 5.9; 2.11: 3.2645 | 3.2566
| 1 | 5 | 2.01: 5; 2.02: < 5; 2.04: 13.8151; 2.05: 3.1; 2.06: 3; 2.07: < 10; 2.08: 3; 2.09: 4; 2.10: 3.3; 2.11: 3.3226 | 78
| 1 | 6 | 2.01: 6; 2.02: < 5; 2.04: 8.5012; 2.05: 3.1; 2.06: 2.9; 2.07: < 10; 2.08: 3; 2.09: 3; 2.10: 4; 2.11: 3.2984
| Mean |  | 2.01: 5.3333; 2.04: 13.0674; 2.05: 3.2333; 2.06: 3.05; 2.08: 3; 2.09: 3.6667; 2.10: 3.7667; 2.11: 3.2497
| Median |  | 2.01: 5; 2.04: 11.1582; 2.05: 3.2; 2.06: 3.05; 2.08: 3; 2.09: 4; 2.10: 3.55; 2.11: 3.2814
| Std Dev. |  | 2.01: 0.5164; 2.04: 6.8634; 2.05: 0.19664; 2.06: 0.10488; 2.08: 0; 2.09: 0.5164; 2.10: 1.1843; 2.11: 0.11703
| Rel.Std.Dev. |  | 2.01: 9.682%; 2.04: 52.52%; 2.05: 6.082%; 2.06: 3.439%; 2.08: 0%; 2.09: 14.08%; 2.10: 31.44%; 2.11: 3.601%
| PDM3 |  | 2.01: 63.77%; 2.04: 301.3%; 2.05: -0.7148%; 2.06: -6.344%; 2.08: -7.88%; 2.09: 12.59%; 2.10: 15.66%; 2.11: -0.2129%
| Z-Score (Absolute) |  | 2.01: 2.92; 2.02: 1.49; 2.04: 14.96; 2.05: 0.35; 2.06: 0.64; 2.07: 2.4; 2.08: 0.71; 2.09: 0.32; 2.10: 0.48; 2.11: 0.32
| Analytical results for Cr in OREAS 466 (Certified Value 170 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 0
0
| 1 | 1 | 197 | 231.946 | 135.156 | 175.633 | 150 | 200 | 200 | 150 | 165 | 155 | 146.955 | [7 empty] | 1
| 1 | 2 | 189 | 207.5775 | 204.889 | 174.0915 | 160 | 200 | 200 | 150 | 165 | 212 | 154.97 | [7 empty] | 0
| 1 | 3 | 184 | 207.7115 | 168.9463 | 174.4233 | 150 | 200 | 100 | 150 | 160 | 170 | 156.73 | [7 empty] | 16
| 1 | 4 | 192 | 210.533 | 136.6142 | 175.8592 | 150 | 200 | 200 | 150 | 190 | 190 | 149.165 | [7 empty] | 170.2198
| 1 | 5 | 199 | 169.9015 | 150.3453 | 170.9338 | 160 | 200 | 200 | 150 | 170 | 147 | 165.735 | [7 empty] | 79
| 1 | 6 | 177 | 212.0245 | 135.1036 | 174.9518 | 150 | 200 | 200 | 150 | 165 | 157 | 153.615
| Mean |  | 189.6667 | 206.6157 | 155.1758 | 174.3154 | 153.3333 | 200 | 183.3333 | 150 | 169.1667 | 171.8333 | 154.5283
| Median |  | 190.5 | 209.1223 | 143.4798 | 174.6875 | 150 | 200 | 200 | 150 | 165 | 163.5 | 154.2925
| Std Dev. |  | 8.2381 | 20.1806 | 27.686 | 1.79 | 5.164 | 0 | 40.8248 | 0 | 10.6849 | 24.7662 | 6.5968
| Rel.Std.Dev. |  | 4.343% | 9.767% | 17.84% | 1.027% | 3.368% | 0% | 22.27% | 0% | 6.316% | 14.41% | 4.269%
| PDM3 |  | 11.42% | 21.38% | -8.838% | 2.406% | -9.92% | 17.5% | 7.704% | -11.88% | -0.6187% | 0.9479% | -9.218%
| Z-Score (Absolute) |  | 0.94 | 1.73 | 0.64 | 0.24 | 0.73 | NA | NA | 0.88 | 0 | 0.12 | 0.67
Results from laboratories 9 and 15 were removed due to their 100 ppm reading resolution.
| Analytical results for Cr2O3 in OREAS 466 (Certified Value 249 ppm) | [19 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 0
0
| 1 | 1 | 287.9 | 338.9891 | 197.5305 | 256.6876 | 219 | 290 | 290 | 219 | 241.1 | 226.53 | 214.7747 | [7 empty] | 1
| 1 | 2 | 276.2 | 303.3745 | 299.4453 | 254.4348 | 234 | 290 | 290 | 219 | 241.1 | 309.84 | 226.4887 | [7 empty] | 0
| 1 | 3 | 268.9 | 303.5704 | 246.915 | 254.9197 | 219 | 290 | 150 | 219 | 233.8 | 248.46 | 229.0609 | [7 empty] | 16
| 1 | 4 | 280.6 | 307.694 | 199.6616 | 257.0182 | 219 | 290 | 290 | 219 | 277.7 | 277.68 | 218.0046 | [7 empty] | 248.7762
| 1 | 5 | 290.8 | 248.311 | 219.7297 | 249.8197 | 234 | 290 | 290 | 219 | 248.5 | 214.84 | 242.2217 | [7 empty] | 80
| 1 | 6 | 258.7 | 309.8738 | 197.454 | 255.692 | 219 | 290 | 290 | 219 | 241.1 | 229.46 | 224.5083
| Mean |  | 277.1833 | 301.9688 | 226.7894 | 254.762 | 224 | 290 | 266.6667 | 219 | 247.2167 | 251.135 | 225.8432
| Median |  | 278.4 | 305.6322 | 209.6957 | 255.3059 | 219 | 290 | 290 | 219 | 241.1 | 238.96 | 225.4985
| Std Dev. |  | 12.0262 | 29.4939 | 40.4631 | 2.6162 | 7.746 | 0 | 57.1548 | 0 | 15.6405 | 36.1956 | 9.6412
| Rel.Std.Dev. |  | 4.339% | 9.767% | 17.84% | 1.027% | 3.458% | 0% | 21.43% | 0% | 6.327% | 14.41% | 4.269%
| PDM3 |  | 11.42% | 21.38% | -8.838% | 2.406% | -9.959% | 16.57% | 7.191% | -11.97% | -0.6269% | 0.9482% | -9.218%
| Z-Score (Absolute) |  | 0.94 | 1.73 | 0.64 | 0.24 | 0.73 | NA | NA | 0.88 | 0 | 0.12 | 0.67
Results from laboratories 9 and 15 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 466 (Certified Value 10.9 ppm) | [19 empty] | Certified
Cs | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.04: 2.07; 2.05: 2.09; 2.06: 2.17; 2.07: 2.23; 2.08: 2.25 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*OES/MS; 2.04: PF*OES/MS; 2.05: BF*MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS | 1
2
| 1 | 1 | 2.01: 12; 2.02: 10.5611; 2.04: 10.3; 2.05: 11.1; 2.06: 10.8; 2.07: 10.6; 2.08: 11.2 | 1
| 1 | 2 | 2.01: 12; 2.02: 10.3984; 2.04: 10.6; 2.05: 11.1; 2.06: 11; 2.07: 11.2; 2.08: 10.4 | 0
| 1 | 3 | 2.01: 17; 2.02: 10.8341; 2.04: 10.3; 2.05: 11.1; 2.06: 11.3; 2.07: 11.3; 2.08: 10.9 | 16
| 1 | 4 | 2.01: 17; 2.02: 10.6397; 2.04: 10.2; 2.05: 11.1; 2.06: 10.8; 2.07: 11.4; 2.08: 10.9 | 10.8718
| 1 | 5 | 2.01: 11; 2.02: 10.226; 2.04: 10.4; 2.05: 11.1; 2.06: 11.5; 2.07: 10.9; 2.08: 10.7 | 81
| 1 | 6 | 2.01: 12; 2.02: 10.3949; 2.04: 9.8; 2.05: 10.8; 2.06: 10.8; 2.07: 10.2; 2.08: 10.9
| Mean |  | 2.01: 13.5; 2.02: 10.509; 2.04: 10.2667; 2.05: 11.05; 2.06: 11.0333; 2.07: 10.9333; 2.08: 10.8333
| Median |  | 2.01: 12; 2.02: 10.4798; 2.04: 10.3; 2.05: 11.1; 2.06: 10.9; 2.07: 11.05; 2.08: 10.9
| Std Dev. |  | 2.01: 2.7386; 2.02: 0.21484; 2.04: 0.26583; 2.05: 0.12247; 2.06: 0.30111; 2.07: 0.46332; 2.08: 0.26583
| Rel.Std.Dev. |  | 2.01: 20.29%; 2.02: 2.044%; 2.04: 2.589%; 2.05: 1.108%; 2.06: 2.729%; 2.07: 4.238%; 2.08: 2.454%
| PDM3 |  | 2.01: 24.17%; 2.02: -3.337%; 2.04: -5.566%; 2.05: 1.639%; 2.06: 1.486%; 2.07: 0.5659%; 2.08: -0.3539%
| Z-Score (Absolute) |  | 2.01: NA; 2.02: 1.59; 2.04: 2.63; 2.05: 0.71; 2.06: 0.64; 2.07: 0.21; 2.08: 0.21
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 466 (Certified Value 35.6 ppm) | [19 empty] | Certified
Cu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.05; 2.05: 2.07; 2.06: 2.09; 2.07: 2.15; 2.08: 2.17; 2.09: 2.23; 2.10: 2.25 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*OES/MS; 2.06: PF*OES; 2.07: PF*OES; 2.08: PF*OES/MS; 2.09: PF*MS; 2.10: BF*MS | 1
2
| 1 | 1 | 2.01: NA; 2.02: 32.8; 2.04: 74.2684; 2.05: 33; 2.06: < 50; 2.07: 100; 2.08: 40; 2.09: 35; 2.10: 35 | 1
| 1 | 2 | 2.01: NA; 2.02: 35.46; 2.04: 95.015; 2.05: 33; 2.06: < 50; 2.07: 70; 2.08: 40; 2.09: 40; 2.10: 34 | 0
| 1 | 3 | 2.01: NA; 2.02: 32.03; 2.04: 102.3735; 2.05: 37; 2.06: < 50; 2.07: 70; 2.08: 40; 2.09: 35; 2.10: 39 | 16
| 1 | 4 | 2.01: NA; 2.02: 36.04; 2.04: 73.143; 2.05: 42; 2.06: < 50; 2.07: 50; 2.08: 40; 2.09: 35; 2.10: 34 | 35.5507
| 1 | 5 | 2.01: NA; 2.02: 33.66; 2.04: 80.0689; 2.05: 34; 2.06: < 50; 2.07: 60; 2.08: 30; 2.09: 35; 2.10: 34 | 82
| 1 | 6 | 2.01: NA; 2.02: 35.73; 2.04: 71.5242; 2.05: 53; 2.06: < 50; 2.07: < 50; 2.08: 30; 2.09: 35; 2.10: 35
| Mean |  | 2.02: 34.2867; 2.04: 82.7322; 2.05: 38.6667; 2.07: 70; 2.08: 36.6667; 2.09: 35.8333; 2.10: 35.1667
| Median |  | 2.02: 34.56; 2.04: 77.1686; 2.05: 35.5; 2.07: 70; 2.08: 40; 2.09: 35; 2.10: 34.5
| Std Dev. |  | 2.02: 1.687; 2.04: 12.907; 2.05: 7.8145; 2.07: 18.7083; 2.08: 5.164; 2.09: 2.0412; 2.10: 1.9408
| Rel.Std.Dev. |  | 2.02: 4.92%; 2.04: 15.6%; 2.05: 20.21%; 2.07: 26.73%; 2.08: 14.08%; 2.09: 5.696%; 2.10: 5.519%
| PDM3 |  | 2.02: -3.555%; 2.04: 132.7%; 2.05: 8.765%; 2.07: 96.9%; 2.08: 3.139%; 2.09: 0.7951%; 2.10: -1.08%
| Z-Score (Absolute) |  | 2.01: NA; 2.02: 0.6; 2.04: 14.31; 2.05: 0.74; 2.06: 3.46; 2.07: 8.08; 2.08: 0.13; 2.09: 0.13; 2.10: 0.33
| Analytical results for Dy in OREAS 466 (Certified Value 7.64 ppm) | [19 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 8.2 | 7 | 7.529 | 8.1 | 8.12 | 7.42 | 8.1 | 7.5 | 7.6 | 7.1 | 7.3 | 8.1 | 6.5 | 16.7 | 7.86 | 7.6812 |  |  | 1
| 1 | 2 | 8 | 7 | 6.9982 | 7.5 | 8.41 | 7.67 | 7 | 7.4 | 8.6 | 7.7 | 7.4 | 8.2 | 7.5 | 16.8 | 7.81 | 8.1428 |  |  | 0
| 1 | 3 | 8 | 7 | 7.5246 | 7.4 | 8.16 | 7.31 | 7.3 | 7.9 | 7.1 | 7.9 | 7.8 | 8 | 7 | 16.9 | 7.57 | 7.8961 |  |  | 16
| 1 | 4 | 7.9 | 7 | 7.4106 | 7.4 | 8.4 | 7.3 | 7.1 | 7.8 | 7.5 | 7.4 | 7.9 | 8 | 7 | 16.8 | 7.65 | 7.6097 |  |  | 7.6412
| 1 | 5 | 7.3 | 7 | 7.4614 | 6.9 | 8.56 | 7.52 | 6.9 | 7.5 | 8.2 | 8.5 | 7.4 | 8.2 | 7.6 | 16.1 | 7.79 | 7.47 |  |  | 83
| 1 | 6 | 7.7 | 7 | 7.3202 | 8 | 8.1 | 7.11 | 7.4 | 7.7 | 7.5 | 7.3 | 7.7 | 8.1 | 6.6 | 16.3 | 7.79 | 7.5683
| Mean |  | 7.85 | 7 | 7.374 | 7.55 | 8.2917 | 7.3883 | 7.3 | 7.6333 | 7.75 | 7.65 | 7.5833 | 8.1 | 7.0333 | 16.6 | 7.745 | 7.728
| Median |  | 7.95 | 7 | 7.436 | 7.45 | 8.28 | 7.365 | 7.2 | 7.6 | 7.55 | 7.55 | 7.55 | 8.1 | 7 | 16.75 | 7.79 | 7.6454
| Std Dev. |  | 0.31464 | 0 | 0.19992 | 0.44159 | 0.19041 | 0.19447 | 0.43359 | 0.19664 | 0.54681 | 0.50498 | 0.24833 | 0.089443 | 0.45019 | 0.32249 | 0.11059 | 0.24844
| Rel.Std.Dev. |  | 4.008% | 0% | 2.711% | 5.849% | 2.296% | 2.632% | 5.94% | 2.576% | 7.056% | 6.601% | 3.275% | 1.104% | 6.401% | 1.943% | 1.428% | 3.215%
| PDM3 |  | 2.732% | -8.392% | -3.497% | -1.194% | 8.512% | -3.309% | -4.465% | -0.1031% | 1.424% | 0.115% | -0.7575% | 6.004% | -7.955% | 117.2% | 1.358% | 1.136%
| Z-Score (Absolute) |  | 0.67 | NA | 0.93 | 0.34 | 2.16 | 0.88 | 1.18 | 0.06 | 0.34 | 0 | 0.22 | 1.52 | 2.08 | 30.18 | 0.32 | 0.26
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for Dy2O3 in OREAS 466 (Certified Value 8.72 ppm) | [19 empty] | Certified
Dy2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 9.41 | 8 | 8.6411 | 9.3 | 9.319 | 8.516 | 9.3 | 8.61 | 8.72 | 8.15 | 8.38 | 9.3 | 7.46 | 19.17 | 9.021 | 8.8157 |  |  | 1
| 1 | 2 | 9.18 | 8 | 8.0318 | 8.61 | 9.652 | 8.803 | 8.03 | 8.49 | 9.87 | 8.84 | 8.49 | 9.41 | 8.608 | 19.28 | 8.964 | 9.3455 |  |  | 0
| 1 | 3 | 9.18 | 8 | 8.636 | 8.49 | 9.365 | 8.39 | 8.38 | 9.07 | 8.15 | 9.07 | 8.95 | 9.18 | 8.034 | 19.4 | 8.688 | 9.0624 |  |  | 16
| 1 | 4 | 9.07 | 8 | 8.5051 | 8.49 | 9.641 | 8.378 | 8.15 | 8.95 | 8.61 | 8.49 | 9.07 | 9.18 | 8.034 | 19.28 | 8.78 | 8.7336 |  |  | 8.7208
| 1 | 5 | 8.38 | 8 | 8.5634 | 7.92 | 9.824 | 8.631 | 7.92 | 8.61 | 9.41 | 9.76 | 8.49 | 9.41 | 8.723 | 18.48 | 8.941 | 8.5734 |  |  | 84
| 1 | 6 | 8.84 | 8 | 8.4015 | 9.18 | 9.296 | 8.16 | 8.49 | 8.84 | 8.61 | 8.38 | 8.84 | 9.3 | 7.575 | 18.71 | 8.941 | 8.6861
| Mean |  | 9.01 | 8 | 8.4631 | 8.665 | 9.5162 | 8.4797 | 8.3783 | 8.7617 | 8.895 | 8.7817 | 8.7033 | 9.2967 | 8.0723 | 19.0533 | 8.8892 | 8.8694
| Median |  | 9.125 | 8 | 8.5343 | 8.55 | 9.503 | 8.453 | 8.265 | 8.725 | 8.665 | 8.665 | 8.665 | 9.3 | 8.034 | 19.225 | 8.941 | 8.7747
| Std Dev. |  | 0.35978 | 0 | 0.22945 | 0.50749 | 0.21863 | 0.22331 | 0.4991 | 0.22649 | 0.62673 | 0.58122 | 0.28619 | 0.10289 | 0.51682 | 0.36963 | 0.12707 | 0.28513
| Rel.Std.Dev. |  | 3.993% | 0% | 2.711% | 5.857% | 2.297% | 2.633% | 5.957% | 2.585% | 7.046% | 6.619% | 3.288% | 1.107% | 6.402% | 1.94% | 1.43% | 3.215%
| PDM3 |  | 3.317% | -8.265% | -2.954% | -0.6394% | 9.121% | -2.765% | -3.927% | 0.4691% | 1.998% | 0.6984% | -0.1998% | 6.604% | -7.435% | 118.5% | 1.931% | 1.705%
| Z-Score (Absolute) |  | 0.61 | 1.87 | 0.78 | 0.27 | 1.9 | 0.74 | 1 | 0.02 | 0.32 | 0.02 | 0.17 | 1.34 | 1.78 | 26.17 | 0.3 | 0.25
| Analytical results for Er in OREAS 466 (Certified Value 3.65 ppm) | [19 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 3.7 | 4 | 3.5713 | 3.8 | 3.04 | 3.61 | 4.5 | 3.8 | 3.6 | 3.3 | 3.7 | 3.7 | 4 | 13.8 | 3.71 | 3.5394 |  |  | 1
| 1 | 2 | 3.8 | 4 | 3.2803 | 3.6 | 3.33 | 3.73 | 3.7 | 3.8 | 3.9 | 3.3 | 3.6 | 3.8 | 4.3 | 13.5 | 3.76 | 3.5529 |  |  | 0
| 1 | 3 | 3.5 | 4 | 3.7782 | 3.8 | 3.15 | 3.54 | 3.8 | 3.9 | 3.2 | 3.6 | 3.7 | 3.6 | 3.9 | 13.7 | 3.69 | 3.7789 |  |  | 16
| 1 | 4 | 3.7 | 4 | 3.5872 | 4.2 | 3.24 | 3.56 | 4.7 | 3.4 | 4 | 3.8 | 3.8 | 3.6 | 4.1 | 13.5 | 3.73 | 3.6905 |  |  | 3.6537
| 1 | 5 | 3.5 | 3 | 3.6659 | 3.8 | 3.01 | 3.58 | 3.8 | 3.7 | 3.9 | 3 | 3.4 | 3.6 | 4.2 | 13 | 3.68 | 3.3116 |  |  | 85
| 1 | 6 | 3.9 | 4 | 3.4812 | 4 | 3.33 | 3.6 | 3.7 | 3.5 | 3.7 | 3.2 | 3.8 | 3.7 | 4.1 | 13.2 | 3.64 | 3.5978
| Mean |  | 3.6833 | 3.8333 | 3.5607 | 3.8667 | 3.1833 | 3.6033 | 4.0333 | 3.6833 | 3.7167 | 3.3667 | 3.6667 | 3.6667 | 4.1 | 13.45 | 3.7017 | 3.5785
| Median |  | 3.7 | 4 | 3.5793 | 3.8 | 3.195 | 3.59 | 3.8 | 3.75 | 3.8 | 3.3 | 3.7 | 3.65 | 4.1 | 13.5 | 3.7 | 3.5754
| Std Dev. |  | 0.16021 | 0.40825 | 0.16978 | 0.20656 | 0.13995 | 0.067132 | 0.44572 | 0.19408 | 0.29269 | 0.28752 | 0.15055 | 0.08165 | 0.14142 | 0.30166 | 0.041673 | 0.15907
| Rel.Std.Dev. |  | 4.35% | 10.65% | 4.768% | 5.342% | 4.396% | 1.863% | 11.05% | 5.269% | 7.875% | 8.54% | 4.106% | 2.227% | 3.449% | 2.243% | 1.126% | 4.445%
| PDM3 |  | 0.8115% | 4.917% | -2.545% | 5.829% | -12.87% | -1.378% | 10.39% | 0.8115% | 1.724% | -7.856% | 0.3553% | 0.3553% | 12.22% | 268.1% | 1.313% | -2.057%
| Z-Score (Absolute) |  | 0 | NA | 0.79 | 1.18 | 3.22 | 0.51 | 2.25 | 0 | 0.21 | 2.04 | 0.11 | 0.11 | 2.68 | 62.83 | 0.12 | 0.67
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for Er2O3 in OREAS 466 (Certified Value 4.24 ppm) | [19 empty] | Certified
Er2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 4.23 | 4.6 | 4.0838 | 4.35 | 3.476 | 4.128 | 5.15 | 4.35 | 4.12 | 3.77 | 4.23 | 4.23 | 4.574 | 15.78 | 4.2424 | 4.0473 |  |  | 1
| 1 | 2 | 4.35 | 4.6 | 3.751 | 4.12 | 3.808 | 4.265 | 4.23 | 4.35 | 4.46 | 3.77 | 4.12 | 4.35 | 4.917 | 15.44 | 4.2996 | 4.0628 |  |  | 0
| 1 | 3 | 4 | 4.6 | 4.3203 | 4.35 | 3.602 | 4.048 | 4.35 | 4.46 | 3.66 | 4.12 | 4.23 | 4.12 | 4.46 | 15.67 | 4.2195 | 4.3212 |  |  | 16
| 1 | 4 | 4.23 | 4.6 | 4.102 | 4.8 | 3.705 | 4.071 | 5.37 | 3.89 | 4.57 | 4.35 | 4.35 | 4.12 | 4.688 | 15.44 | 4.2653 | 4.2201 |  |  | 4.2414
| 1 | 5 | 4 | 3.4 | 4.192 | 4.35 | 3.442 | 4.094 | 4.35 | 4.23 | 4.46 | 3.43 | 3.89 | 4.12 | 4.803 | 14.87 | 4.2081 | 3.7868 |  |  | 86
| 1 | 6 | 4.46 | 4.6 | 3.9808 | 4.57 | 3.808 | 4.117 | 4.23 | 4 | 4.23 | 3.66 | 4.35 | 4.23 | 4.688 | 15.09 | 4.1623 | 4.1141
| Mean |  | 4.2117 | 4.4 | 4.0717 | 4.4233 | 3.6402 | 4.1205 | 4.6133 | 4.2133 | 4.25 | 3.85 | 4.195 | 4.195 | 4.6883 | 15.3817 | 4.2329 | 4.092
| Median |  | 4.23 | 4.6 | 4.0929 | 4.35 | 3.6535 | 4.1055 | 4.35 | 4.29 | 4.345 | 3.77 | 4.23 | 4.175 | 4.688 | 15.44 | 4.2309 | 4.0884
| Std Dev. |  | 0.18498 | 0.4899 | 0.19415 | 0.23304 | 0.16014 | 0.076626 | 0.50855 | 0.22295 | 0.33347 | 0.33112 | 0.17271 | 0.093113 | 0.16164 | 0.3451 | 0.047682 | 0.1819
| Rel.Std.Dev. |  | 4.392% | 11.13% | 4.768% | 5.268% | 4.399% | 1.86% | 11.02% | 5.292% | 7.846% | 8.601% | 4.117% | 2.22% | 3.448% | 2.244% | 1.126% | 4.445%
| PDM3 |  | -0.7003% | 3.74% | -4.001% | 4.29% | -14.17% | -2.85% | 8.77% | -0.661% | 0.2035% | -9.227% | -1.093% | -1.093% | 10.54% | 262.7% | -0.2004% | -3.521%
| Z-Score (Absolute) |  | 0 | 0.89 | 0.73 | 1.09 | 2.96 | 0.47 | 2.08 | 0 | 0.2 | 1.88 | 0.1 | 0.1 | 2.47 | 57.91 | 0.11 | 0.62
| Analytical results for Eu in OREAS 466 (Certified Value 4.42 ppm) | [19 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 4.7 | 4 | 4.7684 | 4.4 | 5.06 | 4.36 | 4.7 | 4 | 4.1 | 4.4 | 4.5 | 4.7 | 4.3 | 5 | 4.58 | 4.5197 |  |  | 1
| 1 | 2 | 4.5 | 4 | 4.5809 | 4.2 | 5.24 | 4.45 | 4.6 | 3.9 | 4.5 | 4.2 | 4.4 | 4.7 | 4.5 | 5.1 | 4.59 | 4.7698 |  |  | 0
| 1 | 3 | 4.7 | 4 | 4.6671 | 4 | 5.26 | 4.25 | 4.3 | 4.2 | 4.4 | 4.5 | 4.5 | 4.5 | 4.2 | 5 | 4.4 | 4.2959 |  |  | 16
| 1 | 4 | 4.5 | 4 | 4.7028 | 4.2 | 5.08 | 4.17 | 4.7 | 4.2 | 4.3 | 4.4 | 4.3 | 4.8 | 4.4 | 5 | 4.43 | 4.4169 |  |  | 4.4164
| 1 | 5 | 4.3 | 4 | 4.5682 | 4.4 | 5.12 | 4.35 | 4.2 | 4.5 | 4.6 | 4.3 | 4.3 | 4.6 | 4.5 | 5 | 4.39 | 4.4136 |  |  | 87
| 1 | 6 | 4.4 | 4 | 4.4743 | 4.1 | 5.05 | 4.16 | 4.3 | 4.4 | 4.3 | 4.3 | 4.6 | 4.9 | 4.2 | 4.9 | 4.33 | 4.24
| Mean |  | 4.5167 | 4 | 4.627 | 4.2167 | 5.135 | 4.29 | 4.4667 | 4.2 | 4.3667 | 4.35 | 4.4333 | 4.7 | 4.35 | 5 | 4.4533 | 4.4427
| Median |  | 4.5 | 4 | 4.624 | 4.2 | 5.1 | 4.3 | 4.45 | 4.2 | 4.35 | 4.35 | 4.45 | 4.7 | 4.35 | 5 | 4.415 | 4.4153
| Std Dev. |  | 0.16021 | 0 | 0.10607 | 0.16021 | 0.092466 | 0.11576 | 0.22509 | 0.22804 | 0.17512 | 0.10488 | 0.12111 | 0.14142 | 0.13784 | 0.063246 | 0.10708 | 0.18819
| Rel.Std.Dev. |  | 3.547% | 0% | 2.292% | 3.799% | 1.801% | 2.698% | 5.039% | 5.429% | 4.01% | 2.411% | 2.732% | 3.009% | 3.169% | 1.265% | 2.405% | 4.236%
| PDM3 |  | 2.271% | -9.428% | 4.768% | -4.522% | 16.27% | -2.862% | 1.139% | -4.9% | -1.126% | -1.503% | 0.3839% | 6.422% | -1.503% | 13.21% | 0.8367% | 0.595%
| Z-Score (Absolute) |  | 0.45 | 2.5 | 1.08 | 1.26 | 3.98 | 0.85 | 0.16 | 1.36 | 0.41 | 0.5 | 0.03 | 1.5 | 0.5 | 3.21 | 0.09 | 0.03
| Analytical results for Eu2O3 in OREAS 466 (Certified Value 5.11 ppm) | [19 empty] | Certified
Eu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 5.44 | 4.6 | 5.5213 | 5.09 | 5.859 | 5.048 | 5.44 | 4.63 | 4.75 | 5.09 | 5.21 | 5.44 | 4.979 | 5.79 | 5.303 | 5.2334 |  |  | 1
| 1 | 2 | 5.21 | 4.6 | 5.3042 | 4.86 | 6.067 | 5.153 | 5.33 | 4.52 | 5.21 | 4.86 | 5.09 | 5.44 | 5.211 | 5.91 | 5.315 | 5.523 |  |  | 0
| 1 | 3 | 5.44 | 4.6 | 5.404 | 4.63 | 6.091 | 4.921 | 4.98 | 4.86 | 5.09 | 5.21 | 5.21 | 5.21 | 4.863 | 5.79 | 5.095 | 4.9742 |  |  | 16
| 1 | 4 | 5.21 | 4.6 | 5.4454 | 4.86 | 5.882 | 4.828 | 5.44 | 4.86 | 4.98 | 5.09 | 4.98 | 5.56 | 5.095 | 5.79 | 5.129 | 5.1143 |  |  | 5.1137
| 1 | 5 | 4.98 | 4.6 | 5.2896 | 5.09 | 5.928 | 5.037 | 4.86 | 5.21 | 5.33 | 4.98 | 4.98 | 5.33 | 5.211 | 5.79 | 5.083 | 5.1105 |  |  | 88
| 1 | 6 | 5.09 | 4.6 | 5.1808 | 4.75 | 5.847 | 4.817 | 4.98 | 5.09 | 4.98 | 4.98 | 5.33 | 5.67 | 4.863 | 5.67 | 5.014 | 4.9095
| Mean |  | 5.2283 | 4.6 | 5.3576 | 4.88 | 5.9457 | 4.9673 | 5.1717 | 4.8617 | 5.0567 | 5.035 | 5.1333 | 5.4417 | 5.037 | 5.79 | 5.1565 | 5.1442
| Median |  | 5.21 | 4.6 | 5.3541 | 4.86 | 5.905 | 4.979 | 5.155 | 4.86 | 5.035 | 5.035 | 5.15 | 5.44 | 5.037 | 5.79 | 5.112 | 5.1124
| Std Dev. |  | 0.18498 | 9.7295e-16 | 0.12281 | 0.18352 | 0.1072 | 0.13416 | 0.26065 | 0.26225 | 0.20255 | 0.12112 | 0.14095 | 0.16266 | 0.15989 | 0.075895 | 0.12396 | 0.21791
| Rel.Std.Dev. |  | 3.538% | 2.115e-14% | 2.292% | 3.761% | 1.803% | 2.701% | 5.04% | 5.394% | 4.006% | 2.406% | 2.746% | 2.989% | 3.174% | 1.311% | 2.404% | 4.236%
| PDM3 |  | 2.241% | -10.05% | 4.768% | -4.571% | 16.27% | -2.863% | 1.133% | -4.929% | -1.116% | -1.54% | 0.3834% | 6.413% | -1.5% | 13.22% | 0.8364% | 0.595%
| Z-Score (Absolute) |  | 0.45 | 2.5 | 1.08 | 1.26 | 3.98 | 0.85 | 0.16 | 1.36 | 0.41 | 0.5 | 0.03 | 1.5 | 0.5 | 3.21 | 0.09 | 0.03
| Analytical results for Fe in OREAS 466 (Certified Value 5.17 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 2
3
| 1 | 1 | 5.181 | 5.0861 | 5.2277 | 5.335 | 5.15 | 4.87 | 5.34 | 5.16 | 5.3 | 4.7043 | 5.097 | [7 empty] | 1
| 1 | 2 | 5.2255 | 5.0541 | 5.4493 | 5.2329 | 5.15 | 4.97 | 5.06 | 5.12 | 5.37 | 4.692 | 5.2034 | [7 empty] | 0
| 1 | 3 | 5.2445 | 5.0768 | 5.1321 | 5.5034 | 5.13 | 4.85 | 5.03 | 5.16 | 5.15 | 4.443 | 5.1108 | [7 empty] | 16
| 1 | 4 | 5.0599 | 5.046 | 5.6994 | 5.3743 | 5.11 | 4.87 | 5.35 | 5.19 | 5.41 | 4.707 | 5.1094 | [7 empty] | 5.1729
| 1 | 5 | 5.2468 | 5.0239 | 5.2706 | 5.1703 | 5.14 | 4.95 | 5.19 | 5.17 | 5.38 | 4.6172 | 5.1533 | [7 empty] | 89
| 1 | 6 | 5.2518 | 5.0418 | 5.328 | 5.3144 | 5.12 | 5.09 | 5.11 | 5.15 | 5.38 | 4.5506 | 5.124
| Mean |  | 5.2016 | 5.0548 | 5.3512 | 5.3217 | 5.1333 | 4.9333 | 5.18 | 5.1583 | 5.3317 | 4.619 | 5.133
| Median |  | 5.235 | 5.05 | 5.2993 | 5.3247 | 5.135 | 4.91 | 5.15 | 5.16 | 5.375 | 4.6546 | 5.1174
| Std Dev. |  | 0.074132 | 0.023117 | 0.2006 | 0.1156 | 0.01633 | 0.090701 | 0.13885 | 0.023166 | 0.096212 | 0.10584 | 0.039497
| Rel.Std.Dev. |  | 1.425% | 0.4573% | 3.749% | 2.172% | 0.3181% | 1.839% | 2.681% | 0.4491% | 1.805% | 2.291% | 0.7695%
| PDM3 |  | 0.554% | -2.284% | 3.446% | 2.876% | -0.7654% | -4.632% | 0.1368% | -0.2821% | 3.069% | -10.71% | -0.7722%
| Z-Score (Absolute) |  | 0.28 | 0.67 | 1.26 | 1.06 | 0.16 | 1.47 | 0.14 | 0 | 1.13 | 3.51 | 0.17
| Analytical results for Fe2O3 in OREAS 466 (Certified Value 7.4 wt.%) | [19 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 2
3
| 1 | 1 | 7.4073 | 7.2717 | 7.474 | 7.6275 | 7.363 | 6.963 | 7.635 | 7.377 | 7.577 | 6.7257 | 7.2872 | [7 empty] | 1
| 1 | 2 | 7.4709 | 7.2258 | 7.7909 | 7.4815 | 7.363 | 7.106 | 7.234 | 7.32 | 7.677 | 6.7082 | 7.4393 | [7 empty] | 0
| 1 | 3 | 7.4981 | 7.2583 | 7.3374 | 7.8682 | 7.334 | 6.934 | 7.191 | 7.377 | 7.363 | 6.3522 | 7.3068 | [7 empty] | 16
| 1 | 4 | 7.2341 | 7.2142 | 8.1485 | 7.6836 | 7.306 | 6.963 | 7.649 | 7.42 | 7.735 | 6.7296 | 7.3049 | [7 empty] | 7.3957
| 1 | 5 | 7.5014 | 7.1826 | 7.5353 | 7.3919 | 7.349 | 7.077 | 7.42 | 7.392 | 7.692 | 6.6012 | 7.3677 | [7 empty] | 90
| 1 | 6 | 7.5085 | 7.2083 | 7.6174 | 7.598 | 7.32 | 7.277 | 7.306 | 7.363 | 7.692 | 6.506 | 7.3258
| Mean |  | 7.4367 | 7.2268 | 7.6506 | 7.6084 | 7.3392 | 7.0533 | 7.4058 | 7.3748 | 7.6227 | 6.6038 | 7.3386
| Median |  | 7.4845 | 7.22 | 7.5763 | 7.6127 | 7.3415 | 7.02 | 7.363 | 7.377 | 7.6845 | 6.6547 | 7.3163
| Std Dev. |  | 0.10599 | 0.033051 | 0.2868 | 0.16528 | 0.023353 | 0.12955 | 0.19874 | 0.033139 | 0.13762 | 0.15131 | 0.056468
| Rel.Std.Dev. |  | 1.425% | 0.4573% | 3.749% | 2.172% | 0.3182% | 1.837% | 2.684% | 0.4493% | 1.805% | 2.291% | 0.7695%
| PDM3 |  | 0.554% | -2.284% | 3.446% | 2.876% | -0.7648% | -4.63% | 0.1366% | -0.2826% | 3.068% | -10.71% | -0.7722%
| Z-Score (Absolute) |  | 0.28 | 0.67 | 1.26 | 1.06 | 0.16 | 1.47 | 0.14 | 0 | 1.13 | 3.51 | 0.17
| Analytical results for Ga in OREAS 466 (Certified Value 22.2 ppm) | [19 empty] | Certified
Ga | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  | 1
| No | No | 2.01: 2.06; 2.02: 2.07; 2.06: 2.09; 2.07: 2.17; 2.08: 2.23; 2.09: 2.24; 2.10: 2.25 |  |  | ppm
|  |  | 2.01: BF*OES/MS; 2.02: PF*OES/MS; 2.06: BF*MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*MS |  |  | 1
2
| 1 | 1 | 2.01: 23.1045; 2.02: 21.8; 2.06: 21.8; 2.07: 23; 2.08: 24; 2.09: 18.3; 2.10: 22 |  |  | 1
| 1 | 2 | 2.01: 22.6066; 2.02: 21.8; 2.06: 21.7; 2.07: 22; 2.08: 25; 2.09: 17.7; 2.10: 21 |  |  | 0
| 1 | 3 | 2.01: 23.0275; 2.02: 21.4; 2.06: 22.1; 2.07: 23; 2.08: 24; 2.09: 18.8; 2.10: 21 |  |  | 16
| 1 | 4 | 2.01: 21.7129; 2.02: 20.3; 2.06: 22.5; 2.07: 23; 2.08: 23; 2.09: 17.8; 2.10: 21 |  |  | 22.2292
| 1 | 5 | 2.01: 21.0485; 2.02: 22.5; 2.06: 21.9; 2.07: 22; 2.08: 24; 2.09: 17.7; 2.10: 20 |  |  | 91
| 1 | 6 | 2.01: 22.6525; 2.02: 22.9; 2.06: 21.4; 2.07: 22; 2.08: 24; 2.09: 18; 2.10: 20
| Mean |  | 2.01: 22.3588; 2.02: 21.7833; 2.06: 21.9; 2.07: 22.5; 2.08: 24; 2.09: 18.05; 2.10: 20.8333
| Median |  | 2.01: 22.6296; 2.02: 21.8; 2.06: 21.85; 2.07: 22.5; 2.08: 24; 2.09: 17.9; 2.10: 21
| Std Dev. |  | 2.01: 0.81055; 2.02: 0.90646; 2.06: 0.37417; 2.07: 0.54772; 2.08: 0.63246; 2.09: 0.43243; 2.10: 0.75277
| Rel.Std.Dev. |  | 2.01: 3.625%; 2.02: 4.161%; 2.06: 1.709%; 2.07: 2.434%; 2.08: 2.635%; 2.09: 2.396%; 2.10: 3.613%
| PDM3 |  | 2.01: 0.5827%; 2.02: -2.006%; 2.06: -1.481%; 2.07: 1.218%; 2.08: 7.966%; 2.09: -18.8%; 2.10: -6.28%
| Z-Score (Absolute) |  | 2.01: 0.52; 2.02: 0.13; 2.06: 0; 2.07: 0.67; 2.08: 2.36; 2.09: 4.33; 2.10: 1.2
| Analytical results for Gd in OREAS 466 (Certified Value 12.1 ppm) | [19 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 1
2
| 1 | 1 | 12.6 | 12 | 11.4443 | 12.6 | 11.94 | 11.55 | 11.8 | 12.1 | 11.6 | 11.6 | 12.5 | 12.5 | 13.9 | 114 | 12.07 | 11.8258 |  |  | 1
| 1 | 2 | 12.3 | 12 | 10.613 | 12.4 | 12.08 | 12.09 | 10.8 | 11.8 | 12.2 | 11.2 | 11.9 | 12.5 | 13.8 | 113.3 | 12.45 | 12.7361 |  |  | 0
| 1 | 3 | 12.4 | 12 | 11.2944 | 13.6 | 12.56 | 11.43 | 11.8 | 11.7 | 11 | 12.2 | 12.6 | 12.5 | 13.8 | 113.6 | 12.4 | 12.9091 |  |  | 16
| 1 | 4 | 12.8 | 12 | 11.2507 | 12 | 12.61 | 11.27 | 11.7 | 12 | 11.1 | 13 | 12.6 | 12.6 | 12.7 | 111.8 | 12.1 | 12.1035 |  |  | 12.1392
| 1 | 5 | 12.3 | 12 | 10.9727 | 12.5 | 12.49 | 11.59 | 11.6 | 11.7 | 11.8 | 12.4 | 12.2 | 12.6 | 12.7 | 108.5 | 12.11 | 11.8378 |  |  | 92
| 1 | 6 | 12.6 | 13 | 11.0547 | 11.8 | 12.83 | 11.56 | 11.8 | 12.1 | 11.4 | 11.8 | 12.7 | 12.6 | 13.8 | 108.8 | 12.3 | 12.3423
| Mean |  | 12.5 | 12.1667 | 11.105 | 12.4833 | 12.4183 | 11.5817 | 11.5833 | 11.9 | 11.5167 | 12.0333 | 12.4167 | 12.55 | 13.45 | 111.6667 | 12.2383 | 12.2925
| Median |  | 12.5 | 12 | 11.1527 | 12.45 | 12.525 | 11.555 | 11.75 | 11.9 | 11.5 | 12 | 12.55 | 12.55 | 13.8 | 112.55 | 12.205 | 12.2229
| Std Dev. |  | 0.2 | 0.40825 | 0.2947 | 0.62743 | 0.33902 | 0.27571 | 0.392 | 0.18974 | 0.44907 | 0.6377 | 0.30605 | 0.054772 | 0.58224 | 2.4541 | 0.16654 | 0.45603
| Rel.Std.Dev. |  | 1.6% | 3.355% | 2.654% | 5.026% | 2.73% | 2.381% | 3.384% | 1.594% | 3.899% | 5.299% | 2.465% | 0.4364% | 4.329% | 2.198% | 1.361% | 3.71%
| PDM3 |  | 2.972% | 0.226% | -8.52% | 2.835% | 2.299% | -4.593% | -4.579% | -1.971% | -5.129% | -0.8724% | 2.285% | 3.384% | 10.8% | 819.9% | 0.8164% | 1.262%
| Z-Score (Absolute) |  | 0.54 | NA | 3.09 | 0.5 | 0.33 | 1.85 | 1.85 | 1.02 | 2.02 | 0.67 | 0.32 | 0.67 | 3.01 | 258.6 | 0.14 | 0
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd2O3 in OREAS 466 (Certified Value 14 ppm) | [19 empty] | Certified
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 1
2
| 1 | 1 | 14.52 | 13.8 | 13.1907 | 14.52 | 13.762 | 13.313 | 13.6 | 13.95 | 13.37 | 13.37 | 14.41 | 14.41 | 16.021 | 131.4 | 13.912 | 13.6304 |  |  | 1
| 1 | 2 | 14.18 | 13.8 | 12.2326 | 14.29 | 13.923 | 13.935 | 12.45 | 13.6 | 14.06 | 12.91 | 13.72 | 14.41 | 15.906 | 130.59 | 14.35 | 14.6797 |  |  | 17
| 1 | 3 | 14.29 | 13.8 | 13.0179 | 15.68 | 14.477 | 13.174 | 13.6 | 13.49 | 12.68 | 14.06 | 14.52 | 14.41 | 15.906 | 130.94 | 14.292 | 14.8791 |  |  | 16
| 1 | 4 | 14.75 | 13.8 | 12.9675 | 13.83 | 14.534 | 12.99 | 13.49 | 13.83 | 12.79 | 14.98 | 14.52 | 14.52 | 14.638 | 128.86 | 13.946 | 13.9505 |  |  | 13.9743
| 1 | 5 | 14.18 | 13.8 | 12.6472 | 14.41 | 14.396 | 13.359 | 13.37 | 13.49 | 13.6 | 14.29 | 14.06 | 14.52 | 14.638 | 125.06 | 13.958 | 13.6443 |  |  | 93
| 1 | 6 | 14.52 | 15 | 12.7416 | 13.6 | 14.788 | 13.324 | 13.6 | 13.95 | 13.14 | 13.6 | 14.64 | 14.52 | 15.906 | 125.4 | 14.177 | 14.2257
| Mean |  | 14.4067 | 14 | 12.7996 | 14.3883 | 14.3133 | 13.3492 | 13.3517 | 13.7183 | 13.2733 | 13.8683 | 14.3117 | 14.465 | 15.5025 | 128.7083 | 14.1058 | 14.1683
| Median |  | 14.405 | 13.8 | 12.8546 | 14.35 | 14.4365 | 13.3185 | 13.545 | 13.715 | 13.255 | 13.83 | 14.465 | 14.465 | 15.906 | 129.725 | 14.0675 | 14.0881
| Std Dev. |  | 0.22801 | 0.4899 | 0.33967 | 0.72513 | 0.39088 | 0.31778 | 0.45111 | 0.21821 | 0.51737 | 0.73336 | 0.35148 | 0.060249 | 0.67112 | 2.83 | 0.192 | 0.52562
| Rel.Std.Dev. |  | 1.583% | 3.499% | 2.654% | 5.04% | 2.731% | 2.381% | 3.379% | 1.591% | 3.898% | 5.288% | 2.456% | 0.4165% | 4.329% | 2.199% | 1.361% | 3.71%
| PDM3 |  | 3.094% | 0.1838% | -8.406% | 2.963% | 2.426% | -4.474% | -4.456% | -1.832% | -5.016% | -0.7584% | 2.414% | 3.511% | 10.94% | 821% | 0.9412% | 1.388%
| Z-Score (Absolute) |  | 0.61 | 0.26 | 3.01 | 0.57 | 0.4 | 1.78 | 1.77 | 0.95 | 1.94 | 0.6 | 0.39 | 0.74 | 3.08 | 258.19 | 0.07 | 0.07
| Analytical results for Ge in OREAS 466 (Indicative Value 2.33 ppm) | [19 empty] | Indicative
Ge | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.07; 2.06: 2.17; 2.07: 2.23; 2.08: 2.25 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS |  |  | 2
2
| 1 | 1 | 2.01: 3; 2.02: 4.2; 2.06: 2; 2.07: 3; 2.08: 2 |  |  | 1
| 1 | 2 | 2.01: 3; 2.02: 3.9; 2.06: 2; 2.07: 3; 2.08: 2 |  |  | 5
| 1 | 3 | 2.01: 3; 2.02: 2.9; 2.06: 2; 2.07: 2; 2.08: 2 |  |  | 16
| 1 | 4 | 2.01: 3; 2.02: 4.4; 2.06: 2; 2.07: 2; 2.08: 2 |  |  | 2.3333
| 1 | 5 | 2.01: 3; 2.02: 3.7; 2.06: 2; 2.07: 2; 2.08: 2 |  |  | 11
| 1 | 6 | 2.01: 3; 2.02: 3.6; 2.06: 2; 2.07: 2; 2.08: 2
| Mean |  | 2.01: 3; 2.02: 3.7833; 2.06: 2; 2.07: 2.3333; 2.08: 2
| Median |  | 2.01: 3; 2.02: 3.8; 2.06: 2; 2.07: 2; 2.08: 2
| Std Dev. |  | 2.01: 0; 2.02: 0.52694; 2.06: 0; 2.07: 0.5164; 2.08: 0
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 13.93%; 2.06: 0%; 2.07: 22.13%; 2.08: 0%
| PDM3 |  | 2.01: 28.57%; 2.02: 62.14%; 2.06: -14.29%; 2.07: 1.554e-13%; 2.08: -14.29%
| Z-Score (Absolute) |  | 2.01: 1.35; 2.02: 2.93; 2.06: 0.67; 2.07: 0; 2.08: 0.67
| Analytical results for Hf in OREAS 466 (Certified Value 7.5 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.24 | 2.25 | 2.26 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | [6 empty] | 2
3
| 1 | 1 | 7.7 | 7 | < 10 | 7.3 | 7 | 8 | 7 | 7 | 8 | 6.8 | 7.72 | 7.491 | [6 empty] | 1
| 1 | 2 | 7.7 | 7 | < 10 | 7.7 | 7 | 8 | 8 | 7 | 7 | 6.5 | 7.95 | 8.0405 | [6 empty] | 0
| 1 | 3 | 8.4 | 7 | < 10 | 7.6 | 7 | 8 | 8 | 7 | 8 | 6.5 | 7.72 | 7.2725 | [6 empty] | 16
| 1 | 4 | 7 | 6 | < 10 | 7.4 | 8 | 8 | 8 | 7 | 8 | 6.7 | 7.6 | 7.4083 | [6 empty] | 7.5001
| 1 | 5 | 7.4 | 7 | < 10 | 7.6 | 7 | 8 | 8 | 8 | 7 | 6.5 | 7.75 | 7.2344 | [6 empty] | 94
| 1 | 6 | 7.5 | 7 | < 10 | 7.6 | 7 | 8 | 7 | 7 | 8 | 6.4 | 7.8 | 7.1192
| Mean |  | 7.6167 | 6.8333 |  | 7.5333 | 7.1667 | 8 | 7.6667 | 7.1667 | 7.6667 | 6.5667 | 7.7567 | 7.4276
| Median |  | 7.6 | 7 |  | 7.6 | 7 | 8 | 8 | 7 | 8 | 6.5 | 7.735 | 7.3404
| Std Dev. |  | 0.46224 | 0.40825 |  | 0.15055 | 0.40825 | 0 | 0.5164 | 0.40825 | 0.5164 | 0.15055 | 0.11535 | 0.32757
| Rel.Std.Dev. |  | 6.069% | 5.974% |  | 1.999% | 5.696% | 0% | 6.736% | 5.696% | 6.736% | 2.293% | 1.487% | 4.41%
| PDM3 |  | 1.554% | -8.89% |  | 0.4432% | -4.446% | 6.665% | 2.221% | -4.446% | 2.221% | -12.45% | 3.421% | -0.9662%
| Z-Score (Absolute) |  | 0.31 | 1.48 | 5.67 | 0.12 | 0.72 | 1.19 | 0.43 | 0.72 | 0.43 | 2.09 | 0.63 | 0.12
| Analytical results for HfO2 in OREAS 466 (Certified Value 8.84 ppm) | [19 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.24 | 2.25 | 2.26 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | [6 empty] | 2
3
| 1 | 1 | 9.08 | 8.3 | < 12 | 8.61 | 8.3 | 9.4 | 8.3 | 8.3 | 9.4 | 8.02 | 9.104 | 8.8341 | [6 empty] | 1
| 1 | 2 | 9.08 | 8.3 | < 12 | 9.08 | 8.3 | 9.4 | 9.4 | 8.3 | 8.3 | 7.67 | 9.375 | 9.4821 | [6 empty] | 9
| 1 | 3 | 9.91 | 8.3 | < 12 | 8.96 | 8.3 | 9.4 | 9.4 | 8.3 | 9.4 | 7.67 | 9.104 | 8.5765 | [6 empty] | 16
| 1 | 4 | 8.26 | 7.1 | < 12 | 8.73 | 9.4 | 9.4 | 9.4 | 8.3 | 9.4 | 7.9 | 8.963 | 8.7366 | [6 empty] | 8.8449
| 1 | 5 | 8.73 | 8.3 | < 12 | 8.96 | 8.3 | 9.4 | 9.4 | 9.4 | 8.3 | 7.67 | 9.14 | 8.5315 | [6 empty] | 95
| 1 | 6 | 8.84 | 8.3 | < 12 | 8.96 | 8.3 | 9.4 | 8.3 | 8.3 | 9.4 | 7.55 | 9.199 | 8.3956
| Mean |  | 8.9833 | 8.1 |  | 8.8833 | 8.4833 | 9.4 | 9.0333 | 8.4833 | 9.0333 | 7.7467 | 9.1475 | 8.7594
| Median |  | 8.96 | 8.3 |  | 8.96 | 8.3 | 9.4 | 9.4 | 8.3 | 9.4 | 7.67 | 9.122 | 8.6565
| Std Dev. |  | 0.54504 | 0.4899 |  | 0.1758 | 0.44907 | 0 | 0.56804 | 0.44907 | 0.56804 | 0.1758 | 0.13586 | 0.3863
| Rel.Std.Dev. |  | 6.067% | 6.048% |  | 1.979% | 5.294% | 0% | 6.288% | 5.294% | 6.288% | 2.269% | 1.485% | 4.41%
| PDM3 |  | 1.566% | -8.421% |  | 0.4349% | -4.087% | 6.276% | 2.131% | -4.087% | 2.131% | -12.42% | 3.422% | -0.9662%
| Z-Score (Absolute) |  | 0.31 | 1.48 | 5.67 | 0.12 | 0.72 | 1.19 | 0.43 | 0.72 | 0.43 | 2.09 | 0.63 | 0.12
| Analytical results for Ho in OREAS 466 (Certified Value 1.37 ppm) | [19 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 1.4 | < 2 | 1.5029 | 1.4 | 1.51 | 1.35 | 1.42 | 1.32 | 1.32 | 1.38 | 1.38 | 1.4 | 1.3 | 2.4 | 1.4 | 1.3382 |  |  | 1
| 1 | 2 | 1.4 | < 2 | 1.3117 | 1.4 | 1.51 | 1.36 | 1.28 | 1.34 | 1.52 | 1.28 | 1.38 | 1.4 | 1.4 | 2.3 | 1.39 | 1.4652 |  |  | 0
| 1 | 3 | 1.4 | < 2 | 1.5155 | 1.4 | 1.64 | 1.33 | 1.28 | 1.48 | 1.26 | 1.24 | 1.4 | 1.3 | 1.3 | 2.4 | 1.36 | 1.3433 |  |  | 16
| 1 | 4 | 1.4 | < 2 | 1.3806 | 1.4 | 1.46 | 1.31 | 1.39 | 1.43 | 1.42 | 1.34 | 1.38 | 1.4 | 1.4 | 2.3 | 1.39 | 1.3726 |  |  | 1.3651
| 1 | 5 | 1.3 | < 2 | 1.4585 | 1.5 | 1.61 | 1.32 | 1.58 | 1.43 | 1.22 | 1.32 | 1.39 | 1.3 | 1.4 | 2.3 | 1.29 | 1.2401 |  |  | 96
| 1 | 6 | 1.3 | < 2 | 1.3815 | 1.3 | 1.5 | 1.31 | 1.23 | 1.46 | 1.32 | 1.3 | 1.36 | 1.4 | 1.4 | 2.4 | 1.36 | 1.4115
| Mean |  | 1.3667 |  | 1.4251 | 1.4 | 1.5383 | 1.33 | 1.3633 | 1.41 | 1.3433 | 1.31 | 1.3817 | 1.3667 | 1.3667 | 2.35 | 1.365 | 1.3618
| Median |  | 1.4 |  | 1.42 | 1.4 | 1.51 | 1.325 | 1.335 | 1.43 | 1.32 | 1.31 | 1.38 | 1.4 | 1.4 | 2.35 | 1.375 | 1.3579
| Std Dev. |  | 0.05164 |  | 0.080089 | 0.063246 | 0.070261 | 0.020976 | 0.12848 | 0.065115 | 0.10985 | 0.04858 | 0.013292 | 0.05164 | 0.05164 | 0.054772 | 0.040373 | 0.076132
| Rel.Std.Dev. |  | 3.779% |  | 5.62% | 4.518% | 4.567% | 1.577% | 9.424% | 4.618% | 8.177% | 3.708% | 0.962% | 3.779% | 3.779% | 2.331% | 2.958% | 5.591%
| PDM3 |  | 0.1113% |  | 4.392% | 2.553% | 12.69% | -2.575% | -0.1329% | 3.286% | -1.598% | -4.04% | 1.21% | 0.1113% | 0.1113% | 72.14% | -0.01077% | -0.2447%
| Z-Score (Absolute) |  | 0 | 8.73 | 1.39 | 0.79 | 4.09 | 0.87 | 0.08 | 1.03 | 0.56 | 1.35 | 0.36 | 0 | 0 | 23.4 | 0.04 | 0.12
| Analytical results for Ho2O3 in OREAS 466 (Certified Value 1.56 ppm) | [19 empty] | Certified
Ho2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 1.6 | < 2.3 | 1.7215 | 1.6 | 1.73 | 1.546 | 1.627 | 1.512 | 1.512 | 1.581 | 1.581 | 1.6 | 1.489 | 2.75 | 1.6037 | 1.5329 |  |  | 1
| 1 | 2 | 1.6 | < 2.3 | 1.5025 | 1.6 | 1.73 | 1.558 | 1.466 | 1.535 | 1.741 | 1.466 | 1.581 | 1.6 | 1.604 | 2.63 | 1.5922 | 1.6784 |  |  | 0
| 1 | 3 | 1.6 | < 2.3 | 1.736 | 1.6 | 1.879 | 1.524 | 1.466 | 1.695 | 1.443 | 1.42 | 1.604 | 1.49 | 1.489 | 2.75 | 1.5579 | 1.5387 |  |  | 16
| 1 | 4 | 1.6 | < 2.3 | 1.5815 | 1.6 | 1.672 | 1.501 | 1.592 | 1.638 | 1.627 | 1.535 | 1.581 | 1.6 | 1.604 | 2.63 | 1.5922 | 1.5723 |  |  | 1.5638
| 1 | 5 | 1.49 | < 2.3 | 1.6707 | 1.72 | 1.844 | 1.512 | 1.81 | 1.638 | 1.398 | 1.512 | 1.592 | 1.49 | 1.604 | 2.63 | 1.4777 | 1.4206 |  |  | 97
| 1 | 6 | 1.49 | < 2.3 | 1.5825 | 1.49 | 1.718 | 1.501 | 1.409 | 1.672 | 1.512 | 1.489 | 1.558 | 1.6 | 1.604 | 2.75 | 1.5579 | 1.6169
| Mean |  | 1.5633 |  | 1.6325 | 1.6017 | 1.7622 | 1.5237 | 1.5617 | 1.615 | 1.5388 | 1.5005 | 1.5828 | 1.5633 | 1.5657 | 2.69 | 1.5636 | 1.5599
| Median |  | 1.6 |  | 1.6266 | 1.6 | 1.73 | 1.518 | 1.529 | 1.638 | 1.512 | 1.5005 | 1.581 | 1.6 | 1.604 | 2.69 | 1.5751 | 1.5555
| Std Dev. |  | 0.056804 |  | 0.091742 | 0.072778 | 0.080619 | 0.023838 | 0.14729 | 0.074453 | 0.12578 | 0.055867 | 0.015224 | 0.056804 | 0.059386 | 0.065727 | 0.046234 | 0.087209
| Rel.Std.Dev. |  | 3.634% |  | 5.62% | 4.544% | 4.575% | 1.565% | 9.431% | 4.61% | 8.174% | 3.723% | 0.9618% | 3.634% | 3.793% | 2.443% | 2.957% | 5.591%
| PDM3 |  | -0.0283% |  | 4.392% | 2.423% | 12.69% | -2.565% | -0.1349% | 3.276% | -1.595% | -4.046% | 1.219% | -0.0283% | 0.1209% | 72.02% | -0.01125% | -0.2447%
| Z-Score (Absolute) |  | 0 | 8.73 | 1.39 | 0.79 | 4.09 | 0.87 | 0.08 | 1.03 | 0.56 | 1.35 | 0.36 | 0 | 0 | 23.4 | 0.04 | 0.12
| Analytical results for In in OREAS 466 (Indicative Value 0.2 ppm) | [19 empty] | Indicative
In | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.06: 2.07; 2.07: 2.17; 2.08: 2.23; 2.09: 2.25 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS |  |  | 2
2
| 1 | 1 | 2.01: 0.21; 2.02: < 0.5; 2.06: < 0.2; 2.07: 0.1; 2.08: < 0.25; 2.09: < 0.2 |  |  | 1
| 1 | 2 | 2.01: 0.201; 2.02: < 0.5; 2.06: < 0.2; 2.07: 0.2; 2.08: < 0.25; 2.09: < 0.2 |  |  | 6
| 1 | 3 | 2.01: 0.24; 2.02: < 0.5; 2.06: < 0.2; 2.07: 0.2; 2.08: < 0.25; 2.09: < 0.2 |  |  | 16
| 1 | 4 | 2.01: 0.225; 2.02: < 0.5; 2.06: < 0.2; 2.07: 0.2; 2.08: < 0.25; 2.09: < 0.2 |  |  | 0.19508
| 1 | 5 | 2.01: 0.203; 2.02: < 0.5; 2.06: < 0.2; 2.07: 0.2; 2.08: < 0.25; 2.09: < 0.2 |  |  | 12
| 1 | 6 | 2.01: 0.262; 2.02: < 0.5; 2.06: < 0.2; 2.07: 0.1; 2.08: < 0.25; 2.09: < 0.2
| Mean |  | 2.01: 0.2235; 2.07: 0.16667
| Median |  | 2.01: 0.2175; 2.07: 0.2
| Std Dev. |  | 2.01: 0.02394; 2.07: 0.05164
| Rel.Std.Dev. |  | 2.01: 10.71%; 2.07: 30.98%
| PDM3 |  | 2.01: 14.57%; 2.07: -14.57%
| Z-Score (Absolute) |  | 2.01: 1.14; 2.02: 1.53; 2.06: 0.67; 2.07: 0.31; 2.08: 0.31; 2.09: 0.67
| Analytical results for K in OREAS 466 (Certified Value 2.41 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [8 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [8 empty] | 2
3
| 1 | 1 | 2.6309 | 2.9342 | 0.53049 | 2.4 | 2.38 | 2.5 | 2.46 | 2.4 | 2.3007 | 2.4226 | [8 empty] | 1
| 1 | 2 | 2.5555 | 2.9262 | 0.52772 | 2.5 | 2.43 | 2.5 | 2.54 | 2.4 | 2.3167 | 2.4399 | [8 empty] | 0
| 1 | 3 | 2.6006 | 2.9252 | 0.66137 | 2.5 | 2.38 | 2.3 | 2.42 | 2.4 | 2.2928 | 2.4462 | [8 empty] | 16
| 1 | 4 | 2.5551 | 2.9021 | 0.54622 | 2.4 | 2.4 | 2.4 | 2.38 | 2.3 | 2.2808 | 2.4446 | [8 empty] | 2.4058
| 1 | 5 | 2.5893 | 2.9318 | 0.55086 | 2.4 | 2.41 | 2.4 | 2.44 | 2.4 | 2.3038 | 2.4231 | [8 empty] | 98
| 1 | 6 | 2.5888 | 2.9284 | 0.4952 | 2.4 | 2.47 | 2.4 | 2.5 | 2.4 | 2.3109 | 2.4515
| Mean |  | 2.5867 | 2.9247 | 0.55198 | 2.4333 | 2.4117 | 2.4167 | 2.4567 | 2.3833 | 2.3009 | 2.438
| Median |  | 2.589 | 2.9273 | 0.53835 | 2.4 | 2.405 | 2.4 | 2.45 | 2.4 | 2.3022 | 2.4423
| Std Dev. |  | 0.02874 | 0.011549 | 0.057053 | 0.05164 | 0.034303 | 0.075277 | 0.057155 | 0.040825 | 0.012866 | 0.012301
| Rel.Std.Dev. |  | 1.111% | 0.3949% | 10.34% | 2.122% | 1.422% | 3.115% | 2.327% | 1.713% | 0.5592% | 0.5046%
| PDM3 |  | 7.519% | 21.57% | -77.06% | 1.144% | 0.2437% | 0.4515% | 2.114% | -0.934% | -4.358% | 1.339%
| Z-Score (Absolute) |  | 2.97 | 9.19 | 34.45 | 0.15 | 0.25 | 0.15 | 0.58 | 0.77 | 2.28 | 0.24
| Analytical results for K2O in OREAS 466 (Certified Value 2.9 wt.%) | [19 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [8 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [8 empty] | 2
3
| 1 | 1 | 3.1692 | 3.5345 | 0.63903 | 2.89 | 2.867 | 3.01 | 2.963 | 2.89 | 2.7714 | 2.9183 | [8 empty] | 1
| 1 | 2 | 3.0784 | 3.5249 | 0.63569 | 3.01 | 2.927 | 3.01 | 3.06 | 2.89 | 2.7907 | 2.9392 | [8 empty] | 0
| 1 | 3 | 3.1327 | 3.5237 | 0.79669 | 3.01 | 2.867 | 2.77 | 2.915 | 2.89 | 2.7619 | 2.9467 | [8 empty] | 16
| 1 | 4 | 3.0779 | 3.4959 | 0.65797 | 2.89 | 2.891 | 2.89 | 2.867 | 2.77 | 2.7475 | 2.9448 | [8 empty] | 2.898
| 1 | 5 | 3.1191 | 3.5317 | 0.66357 | 2.89 | 2.903 | 2.89 | 2.939 | 2.89 | 2.7752 | 2.9189 | [8 empty] | 99
| 1 | 6 | 3.1185 | 3.5275 | 0.59651 | 2.89 | 2.975 | 2.89 | 3.012 | 2.89 | 2.7837 | 2.9531
| Mean |  | 3.1159 | 3.523 | 0.66491 | 2.93 | 2.905 | 2.91 | 2.9593 | 2.87 | 2.7717 | 2.9368
| Median |  | 3.1188 | 3.5262 | 0.6485 | 2.89 | 2.897 | 2.89 | 2.951 | 2.89 | 2.7733 | 2.942
| Std Dev. |  | 0.034619 | 0.013912 | 0.068726 | 0.061968 | 0.041163 | 0.090333 | 0.069029 | 0.04899 | 0.015499 | 0.014818
| Rel.Std.Dev. |  | 1.111% | 0.3949% | 10.34% | 2.115% | 1.417% | 3.104% | 2.333% | 1.707% | 0.5592% | 0.5046%
| PDM3 |  | 7.519% | 21.57% | -77.06% | 1.103% | 0.2405% | 0.413% | 2.115% | -0.9672% | -4.358% | 1.339%
| Z-Score (Absolute) |  | 2.97 | 9.19 | 34.45 | 0.15 | 0.25 | 0.15 | 0.58 | 0.77 | 2.28 | 0.24
| Analytical results for La in OREAS 466 (Certified Value 238 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | BF*MS | 0
0
| 1 | 1 | 243 | 210 | 241.1539 | 253.7 | 272.02 | 230 | 274.5 | 228.4 | 223 | 220 | 225 | 248 | 237 | 246 | 250 | 264.9 | 244.7 | 230.5639 | 1
| 1 | 2 | 237.1 | 210 | 250.9183 | 238.3 | 255.7013 | 224 | 274.09 | 231.2 | 206 | 225 | 245 | 246 | 244 | 245 | 245 | 260.3 | 240.2 | 247.0746 | 0
| 1 | 3 | 249.5 | 209 | 255.7078 | 223.9 | 265.2686 | 226 | 273.62 | 227 | 216 | 222 | 223 | 249 | 239 | 245 | 240 | 255.4 | 235.9 | 233.1657 | 16
| 1 | 4 | 240.4 | 229 | 237.9609 | 256 | 263.7617 | 228 | 273.44 | 224.6 | 225 | 218 | 233 | 245 | 245 | 247 | 240 | 256.7 | 241.4 | 235.1543 | 237.9626
| 1 | 5 | 238 | 205 | 242.9667 | 243.8 | 277.2764 | 220 | 276.51 | 226.6 | 225 | 224 | 237 | 251 | 237 | 251 | 240 | 255.1 | 233.9 | 235.3848 | 100
| 1 | 6 | 235.5 | 218 | 239.1969 | 237 | 268.5137 | 228 | 273.64 | 219.9 | 223 | 228 | 224 | 248 | 242 | 247 | 250 | 257.5 | 231.5 | 232.9983
| Mean |  | 240.5833 | 213.5 | 244.6507 | 242.1167 | 267.0903 | 226 | 274.3 | 226.2833 | 219.6667 | 222.8333 | 231.1667 | 247.8333 | 240.6667 | 246.8333 | 244.1667 | 258.3167 | 237.9333 | 235.7236
| Median |  | 239.2 | 210 | 242.0603 | 241.05 | 266.8911 | 227 | 273.865 | 226.8 | 223 | 223 | 229 | 248 | 240.5 | 246.5 | 242.5 | 257.1 | 238.05 | 234.16
| Std Dev. |  | 5.0996 | 8.6891 | 7.0867 | 11.852 | 7.4053 | 3.5777 | 1.1494 | 3.817 | 7.4744 | 3.6009 | 8.7731 | 2.137 | 3.5024 | 2.2286 | 4.916 | 3.7258 | 4.9947 | 5.8288
| Rel.Std.Dev. |  | 2.12% | 4.07% | 2.897% | 4.895% | 2.773% | 1.583% | 0.419% | 1.687% | 3.403% | 1.616% | 3.795% | 0.8623% | 1.455% | 0.9029% | 2.013% | 1.442% | 2.099% | 2.473%
| PDM3 |  | 1.101% | -10.28% | 2.811% | 1.746% | 12.24% | -5.027% | 15.27% | -4.908% | -7.689% | -6.358% | -2.856% | 4.148% | 1.136% | 3.728% | 2.607% | 8.553% | -0.01231% | -0.9409%
| Z-Score (Absolute) |  | 0 | 2.19 | 0.33 | 0.12 | 2.14 | 1.18 | 2.72 | 1.16 | 1.7 | 1.44 | 0.77 | 0.58 | 0 | 0.5 | 0.29 | 1.43 | 0.22 | 0.4
| Analytical results for La2O3 in OREAS 466 (Certified Value 279 ppm) | [19 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | BF*MS | 0
0
| 1 | 1 | 284.99 | 246.3 | 282.8253 | 297.54 | 319.025 | 269.74 | 321.934 | 267.87 | 261.5 | 258 | 263.9 | 290.9 | 278 | 288.51 | 293.2 | 310.67 | 286.984 | 270.4054 | 1
| 1 | 2 | 278.07 | 246.3 | 294.277 | 279.48 | 299.8865 | 262.71 | 321.453 | 271.15 | 241.6 | 263.9 | 287.3 | 288.5 | 286.2 | 287.34 | 287.3 | 305.28 | 281.707 | 289.7691 | 0
| 1 | 3 | 292.61 | 245.1 | 299.8941 | 262.59 | 311.107 | 265.05 | 320.902 | 266.23 | 253.3 | 260.4 | 261.5 | 292 | 280.3 | 287.34 | 281.5 | 299.53 | 276.664 | 273.4567 | 16
| 1 | 4 | 281.94 | 268.6 | 279.0805 | 300.24 | 309.3397 | 267.4 | 320.69 | 263.41 | 263.9 | 255.7 | 273.3 | 287.3 | 287.3 | 289.68 | 281.5 | 301.06 | 283.114 | 275.789 | 279.0826
| 1 | 5 | 279.13 | 240.4 | 284.9513 | 285.93 | 325.1898 | 258.02 | 324.291 | 265.76 | 263.9 | 262.7 | 278 | 294.4 | 278 | 294.37 | 281.5 | 299.18 | 274.318 | 276.0592 | 101
| 1 | 6 | 276.19 | 255.7 | 280.5301 | 277.95 | 314.9129 | 267.4 | 320.925 | 257.9 | 261.5 | 267.4 | 262.7 | 290.9 | 283.8 | 289.68 | 293.2 | 302 | 271.503 | 273.2604
| Mean |  | 282.155 | 250.4 | 286.9264 | 283.955 | 313.2435 | 265.0533 | 321.6992 | 265.3867 | 257.6167 | 261.35 | 271.1167 | 290.6667 | 282.2667 | 289.4867 | 286.3667 | 302.9533 | 279.0483 | 276.4566
| Median |  | 280.535 | 246.3 | 283.8883 | 282.705 | 313.0099 | 266.225 | 321.189 | 265.995 | 261.5 | 261.55 | 268.6 | 290.9 | 282.05 | 289.095 | 284.4 | 301.53 | 279.1855 | 274.6229
| Std Dev. |  | 5.9801 | 10.2082 | 8.3112 | 13.9011 | 8.6849 | 4.1937 | 1.3481 | 4.4764 | 8.7669 | 4.2184 | 10.2903 | 2.524 | 4.0849 | 2.6112 | 5.75 | 4.3683 | 5.8577 | 6.8361
| Rel.Std.Dev. |  | 2.119% | 4.077% | 2.897% | 4.896% | 2.773% | 1.582% | 0.419% | 1.687% | 3.403% | 1.614% | 3.796% | 0.8684% | 1.447% | 0.902% | 2.008% | 1.442% | 2.099% | 2.473%
| PDM3 |  | 1.101% | -10.28% | 2.811% | 1.746% | 12.24% | -5.027% | 15.27% | -4.907% | -7.692% | -6.354% | -2.854% | 4.151% | 1.141% | 3.728% | 2.61% | 8.553% | -0.01227% | -0.9409%
| Z-Score (Absolute) |  | 0 | 2.19 | 0.33 | 0.12 | 2.14 | 1.18 | 2.72 | 1.16 | 1.7 | 1.44 | 0.77 | 0.58 | 0 | 0.5 | 0.29 | 1.43 | 0.22 | 0.4
| Analytical results for Li in OREAS 466 (Certified Value 44.4 ppm) | [19 empty] | Certified
Li | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.04: 2.04; 2.05: 2.05; 2.06: 2.07; 2.07: 2.09; 2.08: 2.17; 2.09: 2.23; 2.10: 2.25 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*MS; 2.04: PF*OES; 2.05: PF*OES; 2.06: PF*OES/MS; 2.07: PF*OES; 2.08: PF*OES/MS; 2.09: PF*OES; 2.10: PF*MS | 1
2
| 1 | 1 | 2.01: 48; 2.02: 45; 2.04: 45.44; 2.05: 40.3628; 2.06: 44; 2.07: 40; 2.08: 44; 2.09: 45; 2.10: 43 | 1
| 1 | 2 | 2.01: 45; 2.02: 42; 2.04: 47.52; 2.05: 42.5741; 2.06: 45; 2.07: 40; 2.08: 45; 2.09: 45; 2.10: 44 | 0
| 1 | 3 | 2.01: 45; 2.02: 40; 2.04: 46.56; 2.05: 40.4874; 2.06: 46; 2.07: 40; 2.08: 45; 2.09: 45; 2.10: 42 | 16
| 1 | 4 | 2.01: 51; 2.02: 42; 2.04: 45.84; 2.05: 40.5851; 2.06: 45; 2.07: 40; 2.08: 44; 2.09: 40; 2.10: 45 | 44.396
| 1 | 5 | 2.01: 45; 2.02: 40; 2.04: 46.62; 2.05: 40.5687; 2.06: 45; 2.07: 50; 2.08: 44; 2.09: 45; 2.10: 43 | 102
| 1 | 6 | 2.01: 44; 2.02: 45; 2.04: 43.63; 2.05: 41.313; 2.06: 45; 2.07: 50; 2.08: 45; 2.09: 45; 2.10: 45
| Mean |  | 2.01: 46.3333; 2.02: 42.3333; 2.04: 45.935; 2.05: 40.9819; 2.06: 45; 2.07: 43.3333; 2.08: 44.5; 2.09: 44.1667; 2.10: 43.6667
| Median |  | 2.01: 45; 2.02: 42; 2.04: 46.2; 2.05: 40.5769; 2.06: 45; 2.07: 40; 2.08: 44.5; 2.09: 45; 2.10: 43.5
| Std Dev. |  | 2.01: 2.6583; 2.02: 2.2509; 2.04: 1.3371; 2.05: 0.84861; 2.06: 0.63246; 2.07: 5.164; 2.08: 0.54772; 2.09: 2.0412; 2.10: 1.2111
| Rel.Std.Dev. |  | 2.01: 5.737%; 2.02: 5.317%; 2.04: 2.911%; 2.05: 2.071%; 2.06: 1.405%; 2.07: 11.92%; 2.08: 1.231%; 2.09: 4.622%; 2.10: 2.773%
| PDM3 |  | 2.01: 4.364%; 2.02: -4.646%; 2.04: 3.466%; 2.05: -7.69%; 2.06: 1.36%; 2.07: -2.394%; 2.08: 0.2342%; 2.09: -0.5167%; 2.10: -1.643%
| Z-Score (Absolute) |  | 2.01: 1.75; 2.02: 1.48; 2.04: 1.43; 2.05: 2.58; 2.06: 0.67; 2.07: 0.67; 2.08: 0.27; 2.09: 0; 2.10: 0.4
| Analytical results for Lu in OREAS 466 (Certified Value 0.47 ppm) | [19 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.24 | 2.25 | 2.26 |  |  |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS |  |  |  |  | 2
3
| 1 | 1 | 0.5 | < 2 | 0.45759 | 0.65 | 0.46 | 0.46 | 0.51 | 0.4 | 0.49 | 0.5 | 0.49 | 0.6 | 0.46 | < 1 |  |  |  |  | 1
| 1 | 2 | 0.5 | < 2 | 0.46627 | 0.44 | 0.48 | 0.34 | 0.51 | 0.46 | 0.46 | 0.42 | 0.48 | 0.6 | 0.46 | < 1 |  |  |  |  | 0
| 1 | 3 | 0.5 | < 2 | 0.52186 | 0.52 | 0.47 | 0.36 | 0.44 | 0.44 | 0.54 | 0.48 | 0.45 | 0.6 | 0.48 | < 1 |  |  |  |  | 16
| 1 | 4 | 0.5 | < 2 | 0.50451 | 0.42 | 0.46 | 0.35 | 0.51 | 0.44 | 0.46 | 0.53 | 0.45 | 0.6 | 0.48 | < 1 |  |  |  |  | 0.46916
| 1 | 5 | 0.5 | < 2 | 0.46922 | 0.49 | 0.46 | 0.49 | 0.54 | 0.42 | 0.36 | 0.44 | 0.47 | 0.6 | 0.43 | < 1 |  |  |  |  | 103
| 1 | 6 | 0.5 | < 2 | 0.48129 | 0.5 | 0.45 | 0.43 | 0.51 | 0.42 | 0.49 | 0.46 | 0.49 | 0.6 | 0.44 | < 1
| Mean |  | 0.5 |  | 0.48346 | 0.50333 | 0.46333 | 0.405 | 0.50333 | 0.43 | 0.46667 | 0.47167 | 0.47167 | 0.6 | 0.45833
| Median |  | 0.5 |  | 0.47526 | 0.495 | 0.46 | 0.395 | 0.51 | 0.43 | 0.475 | 0.47 | 0.475 | 0.6 | 0.46
| Std Dev. |  | 0 |  | 0.024858 | 0.081158 | 0.010328 | 0.063482 | 0.033267 | 0.020976 | 0.059889 | 0.040208 | 0.018348 | 0 | 0.020412
| Rel.Std.Dev. |  | 0% |  | 5.142% | 16.12% | 2.229% | 15.67% | 6.609% | 4.878% | 12.83% | 8.525% | 3.89% | 0% | 4.454%
| PDM3 |  | 6.573% |  | 3.047% | 7.284% | -1.242% | -13.68% | 7.284% | -8.347% | -0.5319% | 0.5339% | 0.5339% | 27.89% | -2.308%
| Z-Score (Absolute) |  | NA | 12.57 | 0.28 | 0.75 | 0.2 | 1.59 | 0.75 | 0.99 | 0.12 | 0 | 0 | 3.05 | 0.32 | 0.67
Results from laboratory 1 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Lu2O3 in OREAS 466 (Certified Value 0.53 ppm) | [19 empty] | Certified
Lu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.24 | 2.25 | 2.26 |  |  |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS |  |  |  |  | 2
3
| 1 | 1 | 0.57 | < 2.3 | 0.52033 | 0.739 | 0.523 | 0.523 | 0.58 | 0.455 | 0.557 | 0.569 | 0.557 | 0.68 | 0.5231 | < 1.1 |  |  |  |  | 1
| 1 | 2 | 0.57 | < 2.3 | 0.5302 | 0.5 | 0.546 | 0.387 | 0.58 | 0.523 | 0.523 | 0.478 | 0.546 | 0.68 | 0.5231 | < 1.1 |  |  |  |  | 0
| 1 | 3 | 0.57 | < 2.3 | 0.59341 | 0.591 | 0.534 | 0.409 | 0.5 | 0.5 | 0.614 | 0.546 | 0.512 | 0.68 | 0.5458 | < 1.1 |  |  |  |  | 16
| 1 | 4 | 0.57 | < 2.3 | 0.57367 | 0.478 | 0.523 | 0.398 | 0.58 | 0.5 | 0.523 | 0.603 | 0.512 | 0.68 | 0.5458 | < 1.1 |  |  |  |  | 0.53004
| 1 | 5 | 0.57 | < 2.3 | 0.53355 | 0.557 | 0.523 | 0.557 | 0.614 | 0.478 | 0.409 | 0.5 | 0.534 | 0.68 | 0.489 | < 1.1 |  |  |  |  | 104
| 1 | 6 | 0.57 | < 2.3 | 0.54728 | 0.569 | 0.512 | 0.489 | 0.58 | 0.478 | 0.557 | 0.523 | 0.557 | 0.68 | 0.5003 | < 1.1
| Mean |  | 0.57 |  | 0.54974 | 0.57233 | 0.52683 | 0.4605 | 0.57233 | 0.489 | 0.5305 | 0.5365 | 0.53633 | 0.68 | 0.52118
| Median |  | 0.57 |  | 0.54041 | 0.563 | 0.523 | 0.449 | 0.58 | 0.489 | 0.54 | 0.5345 | 0.54 | 0.68 | 0.5231
| Std Dev. |  | 0 |  | 0.028266 | 0.092199 | 0.011686 | 0.072099 | 0.037956 | 0.023647 | 0.068193 | 0.045838 | 0.020675 | 0 | 0.0232
| Rel.Std.Dev. |  | 0% |  | 5.142% | 16.11% | 2.218% | 15.66% | 6.632% | 4.836% | 12.85% | 8.544% | 3.855% | 0% | 4.451%
| PDM3 |  | 7.539% |  | 3.717% | 7.979% | -0.6048% | -13.12% | 7.979% | -7.743% | 0.08693% | 1.219% | 1.187% | 28.29% | -1.671%
| Z-Score (Absolute) |  | 0.67 | 15.7 | 0.18 | 0.77 | 0.43 | 2.18 | 0.77 | 1.43 | 0.33 | 0.18 | 0.18 | 3.68 | 0.58 | 0.67
| Analytical results for Mg in OREAS 466 (Certified Value 0.486 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.4773 | 0.43769 | 0.5202 | 0.51791 | 0.49 | 0.49 | 0.49 | 0.48 | 0.45 | 0.4678 | 0.49526 | [7 empty] | 1
| 1 | 2 | 0.4806 | 0.43825 | 0.48106 | 0.51867 | 0.49 | 0.48 | 0.5 | 0.48 | 0.48 | 0.4644 | 0.50286 | [7 empty] | 0
| 1 | 3 | 0.479 | 0.46558 | 0.54451 | 0.51848 | 0.48 | 0.49 | 0.46 | 0.48 | 0.48 | 0.4648 | 0.49682 | [7 empty] | 16
| 1 | 4 | 0.4986 | 0.44604 | 0.5468 | 0.52175 | 0.48 | 0.48 | 0.49 | 0.48 | 0.45 | 0.4925 | 0.50095 | [7 empty] | 0.48556
| 1 | 5 | 0.4784 | 0.4487 | 0.5275 | 0.52269 | 0.48 | 0.48 | 0.48 | 0.48 | 0.46 | 0.456 | 0.50384 | [7 empty] | 105
| 1 | 6 | 0.4801 | 0.4351 | 0.51484 | 0.51955 | 0.48 | 0.51 | 0.49 | 0.48 | 0.48 | 0.4618 | 0.50045
| Mean |  | 0.48233 | 0.44523 | 0.52248 | 0.51984 | 0.48333 | 0.48833 | 0.485 | 0.48 | 0.46667 | 0.46788 | 0.50003
| Median |  | 0.47955 | 0.44215 | 0.52385 | 0.51911 | 0.48 | 0.485 | 0.49 | 0.48 | 0.47 | 0.4646 | 0.5007
| Std Dev. |  | 0.0080562 | 0.011267 | 0.024005 | 0.0019375 | 0.005164 | 0.01169 | 0.013784 | 0 | 0.015055 | 0.012695 | 0.0033632
| Rel.Std.Dev. |  | 1.67% | 2.531% | 4.594% | 0.3727% | 1.068% | 2.394% | 2.842% | 0% | 3.226% | 2.713% | 0.6726%
| PDM3 |  | -0.664% | -8.306% | 7.605% | 7.061% | -0.4581% | 0.5717% | -0.1148% | -1.145% | -3.891% | -3.64% | 2.981%
| Z-Score (Absolute) |  | 0.04 | 1.66 | 1.71 | 1.59 | 0 | 0.22 | 0.07 | 0.15 | 0.73 | 0.67 | 0.73
| Analytical results for MgO in OREAS 466 (Certified Value 0.805 wt.%) | [19 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.79146 | 0.72578 | 0.8626 | 0.8588 | 0.813 | 0.813 | 0.813 | 0.796 | 0.746 | 0.77571 | 0.82124 | [7 empty] | 1
| 1 | 2 | 0.79693 | 0.72671 | 0.7977 | 0.86006 | 0.813 | 0.796 | 0.829 | 0.796 | 0.796 | 0.77007 | 0.83384 | [7 empty] | 0
| 1 | 3 | 0.79428 | 0.77202 | 0.9029 | 0.85975 | 0.796 | 0.813 | 0.763 | 0.796 | 0.796 | 0.77073 | 0.82383 | [7 empty] | 16
| 1 | 4 | 0.82678 | 0.73963 | 0.9067 | 0.86517 | 0.796 | 0.796 | 0.813 | 0.796 | 0.746 | 0.81666 | 0.83068 | [7 empty] | 0.80515
| 1 | 5 | 0.79328 | 0.74403 | 0.8747 | 0.86672 | 0.796 | 0.796 | 0.796 | 0.796 | 0.763 | 0.75614 | 0.83547 | [7 empty] | 106
| 1 | 6 | 0.7961 | 0.72148 | 0.8537 | 0.86151 | 0.796 | 0.846 | 0.813 | 0.796 | 0.796 | 0.76576 | 0.82984
| Mean |  | 0.7998 | 0.73827 | 0.86638 | 0.862 | 0.80167 | 0.81 | 0.8045 | 0.796 | 0.77383 | 0.77585 | 0.82915
| Median |  | 0.79519 | 0.73317 | 0.86865 | 0.86078 | 0.796 | 0.8045 | 0.813 | 0.796 | 0.7795 | 0.7704 | 0.83026
| Std Dev. |  | 0.01336 | 0.018684 | 0.039801 | 0.0032127 | 0.0087788 | 0.019504 | 0.022854 | 1.2162e-16 | 0.025063 | 0.021049 | 0.0055768
| Rel.Std.Dev. |  | 1.67% | 2.531% | 4.594% | 0.3727% | 1.095% | 2.408% | 2.841% | 1.528e-14% | 3.239% | 2.713% | 0.6726%
| PDM3 |  | -0.664% | -8.306% | 7.605% | 7.061% | -0.4328% | 0.6022% | -0.0809% | -1.137% | -3.89% | -3.64% | 2.981%
| Z-Score (Absolute) |  | 0.04 | 1.66 | 1.71 | 1.59 | 0 | 0.22 | 0.07 | 0.15 | 0.73 | 0.67 | 0.73
| Analytical results for Mn in OREAS 466 (Certified Value 0.01 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.0071 | < 0.01 | 0.0098478 | 0.010444 | 0.0099 | < 0.01 | 0.01 | 0.01 | 0.00953 | < 0.01 | 0.0098 | [7 empty] | 1
| 1 | 2 | 0.0069 | < 0.01 | 0.010826 | 0.010289 | 0.0097 | < 0.01 | 0.01 | 0.011 | 0.00961 | < 0.01 | 0.010027 | [7 empty] | 0
| 1 | 3 | 0.0062 | < 0.01 | 0.012065 | 0.010402 | 0.0099 | < 0.01 | 0.01 | 0.01 | 0.00986 | < 0.01 | 0.009936 | [7 empty] | 16
| 1 | 4 | 0.0096 | < 0.01 | 0.0105 | 0.010629 | 0.0098 | < 0.01 | 0.01 | 0.01 | 0.00987 | < 0.01 | 0.009862 | [7 empty] | 0.010021
| 1 | 5 | 0.0072 | < 0.01 | 0.009913 | 0.010571 | 0.0094 | < 0.01 | 0.01 | 0.01 | 0.00943 | < 0.01 | 0.009978 | [7 empty] | 107
| 1 | 6 | 0.0061 | < 0.01 | 0.0099783 | 0.010483 | 0.0097 | < 0.01 | 0.01 | 0.01 | 0.00953 | < 0.01 | 0.0099415
| Mean |  | 2.01: 0.0071833; 2.04: 0.010522; 2.05: 0.01047; 2.06: 0.0097333; 2.08: 0.01; 2.09: 0.010167; 2.10: 0.0096383; 2.12: 0.0099241
| Median |  | 2.01: 0.007; 2.04: 0.010239; 2.05: 0.010464; 2.06: 0.00975; 2.08: 0.01; 2.09: 0.01; 2.10: 0.00957; 2.12: 0.0099387
| Std Dev. |  | 2.01: 0.0012703; 2.04: 0.00084799; 2.05: 0.00012147; 2.06: 0.00018619; 2.08: 0; 2.09: 0.00040825; 2.10: 0.00018465; 2.12: 8.148e-05
| Rel.Std.Dev. |  | 2.01: 17.68%; 2.04: 8.059%; 2.05: 1.16%; 2.06: 1.913%; 2.08: 0%; 2.09: 4.016%; 2.10: 1.916%; 2.12: 0.821%
| PDM3 |  | 2.01: -28.32%; 2.04: 5%; 2.05: 4.479%; 2.06: -2.868%; 2.08: -0.2067%; 2.09: 1.456%; 2.10: -3.816%; 2.12: -0.9643%
| Z-Score (Absolute) |  | 2.34 | 4.34 | 0.72 | 0.67 | 0 | 4.34 | 0.24 | 0.4 | 0.09 | 4.34 | 0.17
| Analytical results for MnO in OREAS 466 (Certified Value 0.013 wt.%) | [19 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.00917 | < 0.013 | 0.012716 | 0.013485 | 0.01278 | < 0.013 | 0.013 | 0.0129 | 0.012305 | < 0.013 | 0.012654 | [7 empty] | 1
| 1 | 2 | 0.00891 | < 0.013 | 0.013979 | 0.013285 | 0.01252 | < 0.013 | 0.013 | 0.0142 | 0.012409 | < 0.013 | 0.012947 | [7 empty] | 0
| 1 | 3 | 0.00801 | < 0.013 | 0.015579 | 0.013431 | 0.01278 | < 0.013 | 0.013 | 0.0129 | 0.012732 | < 0.013 | 0.01283 | [7 empty] | 16
| 1 | 4 | 0.0124 | < 0.013 | 0.013558 | 0.013724 | 0.01265 | < 0.013 | 0.013 | 0.0129 | 0.012744 | < 0.013 | 0.012734 | [7 empty] | 0.012939
| 1 | 5 | 0.0093 | < 0.013 | 0.0128 | 0.01365 | 0.01214 | < 0.013 | 0.013 | 0.0129 | 0.012176 | < 0.013 | 0.012884 | [7 empty] | 108
| 1 | 6 | 0.00788 | < 0.013 | 0.012884 | 0.013536 | 0.01252 | < 0.013 | 0.013 | 0.0129 | 0.012305 | < 0.013 | 0.012837
| Mean |  | 2.01: 0.0092783; 2.04: 0.013586; 2.05: 0.013519; 2.06: 0.012565; 2.08: 0.013; 2.09: 0.013117; 2.10: 0.012445; 2.12: 0.012814
| Median |  | 2.01: 0.00904; 2.04: 0.013221; 2.05: 0.013511; 2.06: 0.012585; 2.08: 0.013; 2.09: 0.0129; 2.10: 0.012357; 2.12: 0.012833
| Std Dev. |  | 2.01: 0.0016405; 2.04: 0.001095; 2.05: 0.00015685; 2.06: 0.00023847; 2.08: 0; 2.09: 0.00053072; 2.10: 0.00023859; 2.12: 0.00010521
| Rel.Std.Dev. |  | 2.01: 17.68%; 2.04: 8.059%; 2.05: 1.16%; 2.06: 1.898%; 2.08: 0%; 2.09: 4.046%; 2.10: 1.917%; 2.12: 0.821%
| PDM3 |  | 2.01: -28.29%; 2.04: 5%; 2.05: 4.479%; 2.06: -2.891%; 2.08: 0.4713%; 2.09: 1.373%; 2.10: -3.817%; 2.12: -0.9643%
| Z-Score (Absolute) |  | 2.34 | 4.34 | 0.72 | 0.67 | 0 | 4.34 | 0.24 | 0.4 | 0.09 | 4.34 | 0.17
| Analytical results for Mo in OREAS 466 (Certified Value 6.72 ppm) | [19 empty] | Certified
Mo | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.05; 2.04: 2.07; 2.05: 2.17; 2.06: 2.23; 2.07: 2.25 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: BF*MS | 2
3
| 1 | 1 | 2.01: 6; 2.02: 11.4635; 2.04: 6; 2.05: 6; 2.06: 7.4; 2.07: 7 | 1
| 1 | 2 | 2.01: 7; 2.02: 10.211; 2.04: 7; 2.05: 6; 2.06: 7.6; 2.07: 7 | 0
| 1 | 3 | 2.01: 6; 2.02: 11.4688; 2.04: 6; 2.05: 6; 2.06: 7; 2.07: 7 | 16
| 1 | 4 | 2.01: 7; 2.02: 9.2817; 2.04: 7; 2.05: 7; 2.06: 7.2; 2.07: 7 | 6.7233
| 1 | 5 | 2.01: 6; 2.02: 9.6503; 2.04: 7; 2.05: 6; 2.06: 7.3; 2.07: 7 | 109
| 1 | 6 | 2.01: 7; 2.02: 9.5902; 2.04: 7; 2.05: 6; 2.06: 7.2; 2.07: 7
| Mean |  | 2.01: 6.5; 2.02: 10.2776; 2.04: 6.6667; 2.05: 6.1667; 2.06: 7.2833; 2.07: 7
| Median |  | 2.01: 6.5; 2.02: 9.9306; 2.04: 7; 2.05: 6; 2.06: 7.25; 2.07: 7
| Std Dev. |  | 2.01: 0.54772; 2.02: 0.96826; 2.04: 0.5164; 2.05: 0.40825; 2.06: 0.20412; 2.07: 0
| Rel.Std.Dev. |  | 2.01: 8.427%; 2.02: 9.421%; 2.04: 7.746%; 2.05: 6.62%; 2.06: 2.803%; 2.07: 0%
| PDM3 |  | 2.01: -3.322%; 2.02: 52.86%; 2.04: -0.8428%; 2.05: -8.28%; 2.06: 8.329%; 2.07: 4.115%
| Z-Score (Absolute) |  | 2.01: 0.57; 2.02: 5.93; 2.04: 0.29; 2.05: 1.15; 2.06: 0.77; 2.07: 0.29
| Analytical results for Na in OREAS 466 (Indicative Value 0.197 wt.%) | [19 empty] | Indicative
Na | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.26 | wt.%
|  |  | 2.01: BF*OES | 3
3
| 1 | 1 | 2.01: 0.19531 | 1
| 1 | 2 | 2.01: 0.1966 | 7
| 1 | 3 | 2.01: 0.19974 | 16
| 1 | 4 | 2.01: 0.19836 | 0.19733
| 1 | 5 | 2.01: 0.19663 | 13
| 1 | 6 | 2.01: 0.19734
| Mean |  | 2.01: 0.19733
| Median |  | 2.01: 0.19699
| Std Dev. |  | 2.01: 0.0015479
| Rel.Std.Dev. |  | 2.01: 0.7844%
PDM3 | bar: 2.01
Z-Score (Absolute) | bar: 2.01
| Analytical results for Na2O in OREAS 466 (Indicative Value 0.266 wt.%) | [19 empty] | Indicative
Na2O | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.26 | wt.%
|  |  | 2.01: BF*OES | 3
3
| 1 | 1 | 2.01: 0.26328 | 1
| 1 | 2 | 2.01: 0.26502 | 7
| 1 | 3 | 2.01: 0.26925 | 16
| 1 | 4 | 2.01: 0.26739 | 0.266
| 1 | 5 | 2.01: 0.26505 | 13
| 1 | 6 | 2.01: 0.26602
| Mean |  | 2.01: 0.266
| Median |  | 2.01: 0.26554
| Std Dev. |  | 2.01: 0.0020865
| Rel.Std.Dev. |  | 2.01: 0.7844%
| PDM3 |  | 2.01: -1.253e-07%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Nb in OREAS 466 (Certified Value 113 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | 0
0
| 1 | 1 | 113 | 97 | 117.1854 | 101.8635 | 118.7892 | 109.2 | 143.88 | 115.8 | 107 | 115 | 114 | 117 | 119 | NA | 135 | 112 | 110.93 | 109.1065 | 1
| 1 | 2 | 112 | 96 | 118.2509 | 109.3001 | 114.0084 | 106 | 141.13 | 115.2 | 101 | 118 | 124 | 117 | 121 | NA | 135 | 106.6 | 111.7 | 116.8706 | 0
| 1 | 3 | 108 | 99 | 119.72 | 124.9081 | 120.7731 | 107.8 | 140.48 | 114.7 | 101 | 118 | 115 | 117 | 116 | NA | 130 | 107.4 | 112.08 | 110.6061 | 16
| 1 | 4 | 109 | 93 | 118.8259 | 103.1752 | 119.7051 | 104.9 | 141.86 | 115.5 | 109 | 112 | 119 | 116 | 120 | NA | 125 | 104.8 | 111.03 | 111.6527 | 112.9384
| 1 | 5 | 110 | 96 | 118.359 | 112.7335 | 116.7828 | 107.7 | 140.14 | 115.4 | 106 | 117 | 120 | 115 | 120 | NA | 130 | 102.3 | 109.75 | 110.1747 | 110
| 1 | 6 | 111 | 101 | 119.3031 | 101.0231 | 121.9748 | 104.8 | 142.19 | 110.5 | 105 | 122 | 113 | 116 | 121 | NA | 125 | 100.3 | 111.05 | 110.2924
| Mean |  | 110.5 | 97 | 118.6074 | 108.8339 | 118.6722 | 106.7333 | 141.6133 | 114.5167 | 104.8333 | 117 | 117.5 | 116.3333 | 119.5 |  | 130 | 105.5667 | 111.09 | 111.4505
| Median |  | 110.5 | 96.5 | 118.5924 | 106.2377 | 119.2471 | 106.85 | 141.495 | 115.3 | 105.5 | 117.5 | 117 | 116.5 | 120 |  | 130 | 105.7 | 111.04 | 110.4492
| Std Dev. |  | 1.8708 | 2.7568 | 0.89223 | 9.115 | 2.8887 | 1.7773 | 1.3579 | 2.0014 | 3.2506 | 3.3466 | 4.2308 | 0.8165 | 1.8708 |  | 4.4721 | 4.1186 | 0.79732 | 2.7783
| Rel.Std.Dev. |  | 1.693% | 2.842% | 0.7523% | 8.375% | 2.434% | 1.665% | 0.9589% | 1.748% | 3.101% | 2.86% | 3.601% | 0.7019% | 1.566% |  | 3.44% | 3.901% | 0.7177% | 2.493%
| PDM3 |  | -2.159% | -14.11% | 5.02% | -3.634% | 5.077% | -5.494% | 25.39% | 1.397% | -7.177% | 3.596% | 4.039% | 3.006% | 5.81% |  | 15.11% | -6.527% | -1.637% | -1.317%
| Z-Score (Absolute) |  | 0.65 | 2.84 | 0.66 | 0.92 | 0.67 | 1.26 | 4.4 | 0 | 1.57 | 0.4 | 0.48 | 0.29 | 0.81 | NA | 2.51 | 1.45 | 0.56 | 0.5
| Analytical results for Nd in OREAS 466 (Certified Value 155 ppm) | [19 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | 0
0
| 1 | 1 | 157.5 | 147 | 176.641 | 164.7 | 158.0839 | 162 | 169.2 | 145.3 | 150 | 145 | 145.5 | 158.5 | 156 | 160 | 140 | 161.33 | 157.8 | 145.6425 | 1
| 1 | 2 | 153.8 | 148 | 173.3825 | 155.1 | 151.8051 | 161 | 173.89 | 148.3 | 138.5 | 150.5 | 160.5 | 156 | 158.5 | 161 | 139 | 161.74 | 155.7 | 159.5147 | 0
| 1 | 3 | 156 | 149 | 172.5025 | 145.9 | 156.6825 | 158 | 169.49 | 145.7 | 142.5 | 149 | 145 | 159 | 158 | 161 | 140 | 167.97 | 154 | 147.4269 | 16
| 1 | 4 | 152.8 | 151 | 173.74 | 168.7 | 152.3354 | 153 | 171.42 | 143.6 | 147.5 | 143.5 | 151.5 | 154 | 159 | 160 | 136 | 169.44 | 154.9 | 145.7492 | 155.156
| 1 | 5 | 149.6 | 149 | 175.996 | 154.6 | 159.0891 | 154 | 170.1 | 144.9 | 150.5 | 147 | 156.5 | 159 | 156.5 | 163 | 136 | 169.04 | 150.2 | 145.1303 | 111
| 1 | 6 | 152.7 | 151 | 174.0885 | 155.2 | 158.5979 | 160 | 169.03 | 142.5 | 145.5 | 150 | 150.5 | 159.5 | 159.5 | 161 | 138 | 169.25 | 150 | 148.5381
| Mean |  | 153.7333 | 149.1667 | 174.3918 | 157.3667 | 156.099 | 158 | 170.5217 | 145.05 | 145.75 | 147.5 | 151.5833 | 157.6667 | 157.9167 | 161 | 138.1667 | 166.4617 | 153.7667 | 148.6669
| Median |  | 153.3 | 149 | 173.9143 | 155.15 | 157.3832 | 159 | 169.795 | 145.1 | 146.5 | 148 | 151 | 158.75 | 158.25 | 161 | 138.5 | 168.505 | 154.45 | 146.588
| Std Dev. |  | 2.7696 | 1.6021 | 1.5961 | 8.1387 | 3.227 | 3.7417 | 1.8637 | 1.9837 | 4.6233 | 2.8284 | 6.0862 | 2.1833 | 1.3934 | 1.0954 | 1.8348 | 3.8522 | 3.1066 | 5.4662
| Rel.Std.Dev. |  | 1.802% | 1.074% | 0.9152% | 5.172% | 2.067% | 2.368% | 1.093% | 1.368% | 3.172% | 1.918% | 4.015% | 1.385% | 0.8824% | 0.6804% | 1.328% | 2.314% | 2.02% | 3.677%
| PDM3 |  | -0.9169% | -3.86% | 12.4% | 1.425% | 0.6078% | 1.833% | 9.903% | -6.513% | -6.062% | -4.934% | -2.303% | 1.618% | 1.779% | 3.767% | -10.95% | 7.287% | -0.8954% | -4.182%
| Z-Score (Absolute) |  | 0.14 | 0.66 | 2.22 | 0.28 | 0.13 | 0.35 | 1.78 | 1.13 | 1.05 | 0.85 | 0.38 | 0.31 | 0.34 | 0.69 | 1.91 | 1.31 | 0.13 | 0.71
| Analytical results for Nd2O3 in OREAS 466 (Certified Value 181 ppm) | [19 empty] | Certified
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | 0
0
| 1 | 1 | 183.71 | 171.5 | 206.0341 | 192.11 | 184.389 | 188.96 | 197.355 | 169.48 | 174.96 | 169.13 | 169.71 | 184.87 | 181.96 | 186.62 | 163.296 | 188.175 | 184.058 | 169.8775 | 1
| 1 | 2 | 179.39 | 172.6 | 202.2334 | 180.91 | 177.0654 | 187.79 | 202.825 | 172.98 | 161.55 | 175.54 | 187.21 | 181.96 | 184.87 | 187.79 | 162.13 | 188.654 | 181.608 | 186.0579 | 11
| 1 | 3 | 181.96 | 173.8 | 201.2069 | 170.18 | 182.7544 | 184.29 | 197.693 | 169.94 | 166.21 | 173.79 | 169.13 | 185.46 | 184.29 | 187.79 | 163.296 | 195.92 | 179.626 | 171.9587 | 16
| 1 | 4 | 178.23 | 176.1 | 202.6503 | 196.77 | 177.684 | 178.46 | 199.944 | 167.5 | 172.04 | 167.38 | 176.71 | 179.63 | 185.46 | 186.62 | 158.63 | 197.635 | 180.675 | 170.0018 | 180.9739
| 1 | 5 | 174.49 | 173.8 | 205.2817 | 180.33 | 185.5615 | 179.63 | 198.405 | 169.01 | 175.54 | 171.46 | 182.54 | 185.46 | 182.54 | 190.12 | 158.63 | 197.168 | 175.193 | 169.28 | 112
| 1 | 6 | 178.11 | 176.1 | 203.0568 | 181.03 | 184.9886 | 186.62 | 197.157 | 166.21 | 169.71 | 174.96 | 175.54 | 186.04 | 186.04 | 187.79 | 160.963 | 197.413 | 174.96 | 173.2548
| Mean |  | 179.315 | 173.9833 | 203.4105 | 183.555 | 182.0738 | 184.2917 | 198.8965 | 169.1867 | 170.0017 | 172.0433 | 176.8067 | 183.9033 | 184.1933 | 187.7883 | 161.1575 | 194.1608 | 179.3533 | 173.4051
| Median |  | 178.81 | 173.8 | 202.8536 | 180.97 | 183.5717 | 185.455 | 198.049 | 169.245 | 170.875 | 172.625 | 176.125 | 185.165 | 184.58 | 187.79 | 161.5465 | 196.544 | 180.1505 | 170.9803
| Std Dev. |  | 3.232 | 1.8498 | 1.8617 | 9.492 | 3.764 | 4.3633 | 2.1736 | 2.3143 | 5.3908 | 3.2972 | 7.0994 | 2.545 | 1.625 | 1.278 | 2.1404 | 4.4931 | 3.6235 | 6.3758
| Rel.Std.Dev. |  | 1.802% | 1.063% | 0.9152% | 5.171% | 2.067% | 2.368% | 1.093% | 1.368% | 3.171% | 1.917% | 4.015% | 1.384% | 0.8822% | 0.6806% | 1.328% | 2.314% | 2.02% | 3.677%
| PDM3 |  | -0.9167% | -3.863% | 12.4% | 1.426% | 0.6078% | 1.833% | 9.903% | -6.513% | -6.063% | -4.935% | -2.303% | 1.619% | 1.779% | 3.765% | -10.95% | 7.287% | -0.8955% | -4.182%
| Z-Score (Absolute) |  | 0.14 | 0.66 | 2.22 | 0.28 | 0.13 | 0.35 | 1.78 | 1.13 | 1.05 | 0.85 | 0.38 | 0.31 | 0.34 | 0.69 | 1.91 | 1.31 | 0.13 | 0.71
| Analytical results for Ni in OREAS 466 (Certified Value 24.6 ppm) | [19 empty] | Certified
Ni | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.05; 2.04: 2.07; 2.05: 2.09; 2.06: 2.15; 2.07: 2.17; 2.08: 2.23; 2.09: 2.25 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*OES; 2.06: PF*OES; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: BF*MS | 1
1
| 1 | 1 | 2.01: 21; 2.02: 46.6407; 2.04: 20; 2.05: < 50; 2.06: < 50; 2.07: 30; 2.08: 30; 2.09: 12 | 1
| 1 | 2 | 2.01: 29; 2.02: 68.7677; 2.04: 30; 2.05: < 50; 2.06: < 50; 2.07: 20; 2.08: 30; 2.09: 19 | 0
| 1 | 3 | 2.01: 26; 2.02: 73.9637; 2.04: 20; 2.05: < 50; 2.06: < 50; 2.07: 30; 2.08: 25; 2.09: 11 | 16
| 1 | 4 | 2.01: 35; 2.02: 52.8135; 2.04: 30; 2.05: < 50; 2.06: < 50; 2.07: 30; 2.08: 25; 2.09: 20 | 24.6333
| 1 | 5 | 2.01: 30; 2.02: 48.3782; 2.04: 30; 2.05: < 50; 2.06: < 50; 2.07: 30; 2.08: 25; 2.09: 11 | 113
| 1 | 6 | 2.01: 28; 2.02: 42.7308; 2.04: 20; 2.05: < 50; 2.06: < 50; 2.07: 30; 2.08: 25; 2.09: 17
| Mean |  | 2.01: 28.1667; 2.02: 55.5491; 2.04: 25; 2.07: 28.3333; 2.08: 26.6667; 2.09: 15
| Median |  | 2.01: 28.5; 2.02: 50.5959; 2.04: 25; 2.07: 30; 2.08: 25; 2.09: 14.5
| Std Dev. |  | 2.01: 4.6224; 2.02: 12.7781; 2.04: 5.4772; 2.07: 4.0825; 2.08: 2.582; 2.09: 4.1473
| Rel.Std.Dev. |  | 2.01: 16.41%; 2.02: 23%; 2.04: 21.91%; 2.07: 14.41%; 2.08: 9.682%; 2.09: 27.65%
| PDM3 |  | 2.01: 14.34%; 2.02: 125.5%; 2.04: 1.488%; 2.07: 15.02%; 2.08: 8.254%; 2.09: -39.11%
| Z-Score (Absolute) |  | 2.01: 0.99; 2.02: 12.66; 2.04: 0.35; 2.05: 0.35; 2.06: 0.35; 2.07: 1.06; 2.08: 0.35; 2.09: 4.61
| Analytical results for P in OREAS 466 (Certified Value 0.056 wt.%) | [19 empty] | Certified
P | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.04: 2.04; 2.05: 2.05; 2.06: 2.17; 2.07: 2.23; 2.08: 2.25; 2.09: 2.26 | wt.%
|  |  | 2.01: PF*OES; 2.02: PF*MS; 2.04: PF*OES; 2.05: PF*OES; 2.06: PF*OES/MS; 2.07: PF*OES; 2.08: PF*MS; 2.09: BF*OES | 3
3
| 1 | 1 | 2.01: 0.06; 2.02: 0.059; 2.04: 0.052215; 2.05: 0.10197; 2.06: 0.06; 2.07: 0.05; 2.08: 0.054; 2.09: 0.056356 | 1
| 1 | 2 | 2.01: 0.06; 2.02: 0.0585; 2.04: 0.052327; 2.05: 0.089914; 2.06: 0.05; 2.07: 0.046; 2.08: 0.0542; 2.09: 0.064033 | 0
| 1 | 3 | 2.01: 0.06; 2.02: 0.0551; 2.04: 0.050983; 2.05: 0.066136; 2.06: 0.06; 2.07: 0.055; 2.08: 0.055; 2.09: 0.062484 | 16
| 1 | 4 | 2.01: 0.06; 2.02: 0.06; 2.04: 0.051298; 2.05: 0.093011; 2.06: 0.06; 2.07: 0.05; 2.08: 0.0573; 2.09: 0.058555 | 0.056036
| 1 | 5 | 2.01: 0.06; 2.02: 0.0584; 2.04: 0.05135; 2.05: 0.047335; 2.06: 0.05; 2.07: 0.05; 2.08: 0.0532; 2.09: 0.063506 | 114
| 1 | 6 | 2.01: 0.06; 2.02: 0.0618; 2.04: 0.052708; 2.05: 0.071389; 2.06: 0.05; 2.07: 0.057; 2.08: 0.0536; 2.09: 0.059588
| Mean |  | 2.01: 0.06; 2.02: 0.0588; 2.04: 0.051813; 2.05: 0.078292; 2.06: 0.055; 2.07: 0.051333; 2.08: 0.05455; 2.09: 0.060754
| Median |  | 2.01: 0.06; 2.02: 0.05875; 2.04: 0.051783; 2.05: 0.080652; 2.06: 0.055; 2.07: 0.05; 2.08: 0.0541; 2.09: 0.061036
| Std Dev. |  | 2.01: 0; 2.02: 0.00221; 2.04: 0.00069214; 2.05: 0.020328; 2.06: 0.0054772; 2.07: 0.0039833; 2.08: 0.0014775; 2.09: 0.0030613
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 3.758%; 2.04: 1.336%; 2.05: 25.96%; 2.06: 9.959%; 2.07: 7.76%; 2.08: 2.709%; 2.09: 5.039%
| PDM3 |  | 2.01: 7.074%; 2.02: 4.933%; 2.04: -7.535%; 2.05: 39.72%; 2.06: -1.848%; 2.07: -8.392%; 2.08: -2.652%; 2.09: 8.42%
| Z-Score (Absolute) |  | 2.01: 0.6; 2.02: 0.37; 2.04: 0.99; 2.05: 4.15; 2.06: 0.37; 2.07: 1.08; 2.08: 0.46; 2.09: 0.75
| Analytical results for P2O5 in OREAS 466 (Certified Value 0.128 wt.%) | [19 empty] | Certified
P2O5 | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.04: 2.04; 2.05: 2.05; 2.06: 2.17; 2.07: 2.23; 2.08: 2.25; 2.09: 2.26 | wt.%
|  |  | 2.01: PF*OES; 2.02: PF*MS; 2.04: PF*OES; 2.05: PF*OES; 2.06: PF*OES/MS; 2.07: PF*OES; 2.08: PF*MS; 2.09: BF*OES | 3
3
| 1 | 1 | 2.01: 0.137; 2.02: 0.1352; 2.04: 0.11966; 2.05: 0.23367; 2.06: 0.137; 2.07: 0.1146; 2.08: 0.12375; 2.09: 0.12915 | 1
| 1 | 2 | 2.01: 0.137; 2.02: 0.13406; 2.04: 0.11991; 2.05: 0.20605; 2.06: 0.115; 2.07: 0.1054; 2.08: 0.1242; 2.09: 0.14674 | 0
| 1 | 3 | 2.01: 0.137; 2.02: 0.12627; 2.04: 0.11683; 2.05: 0.15156; 2.06: 0.137; 2.07: 0.126; 2.08: 0.12604; 2.09: 0.14319 | 16
| 1 | 4 | 2.01: 0.137; 2.02: 0.1375; 2.04: 0.11755; 2.05: 0.21314; 2.06: 0.137; 2.07: 0.1146; 2.08: 0.13131; 2.09: 0.13418 | 0.12841
| 1 | 5 | 2.01: 0.137; 2.02: 0.13383; 2.04: 0.11767; 2.05: 0.10847; 2.06: 0.115; 2.07: 0.1146; 2.08: 0.12191; 2.09: 0.14553 | 115
| 1 | 6 | 2.01: 0.137; 2.02: 0.14162; 2.04: 0.12079; 2.05: 0.1636; 2.06: 0.115; 2.07: 0.1306; 2.08: 0.12283; 2.09: 0.13655
| Mean |  | 2.01: 0.137; 2.02: 0.13475; 2.04: 0.11874; 2.05: 0.17941; 2.06: 0.126; 2.07: 0.11763; 2.08: 0.12501; 2.09: 0.13922
| Median |  | 2.01: 0.137; 2.02: 0.13463; 2.04: 0.11866; 2.05: 0.18482; 2.06: 0.126; 2.07: 0.1146; 2.08: 0.12398; 2.09: 0.13987
| Std Dev. |  | 2.01: 0; 2.02: 0.0050635; 2.04: 0.0015861; 2.05: 0.046583; 2.06: 0.01205; 2.07: 0.0091147; 2.08: 0.003386; 2.09: 0.0070152
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 3.758%; 2.04: 1.336%; 2.05: 25.96%; 2.06: 9.563%; 2.07: 7.748%; 2.08: 2.709%; 2.09: 5.039%
| PDM3 |  | 2.01: 6.688%; 2.02: 4.933%; 2.04: -7.535%; 2.05: 39.72%; 2.06: -1.878%; 2.07: -8.394%; 2.08: -2.651%; 2.09: 8.42%
| Z-Score (Absolute) |  | 2.01: 0.6; 2.02: 0.37; 2.04: 0.99; 2.05: 4.15; 2.06: 0.37; 2.07: 1.08; 2.08: 0.46; 2.09: 0.75
| Analytical results for Pb in OREAS 466 (Indicative Value 25.7 ppm) | [19 empty] | Indicative
Pb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.05; 2.02: 2.07; 2.04: 2.09; 2.05: 2.15; 2.06: 2.17; 2.07: 2.23; 2.08: 2.25 | ppm
|  |  | 2.01: PF*OES; 2.02: PF*OES/MS; 2.04: PF*OES; 2.05: PF*OES; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS | 1
1
| 1 | 1 | 2.01: 60.0412; 2.02: 22.2; 2.04: < 300; 2.05: < 100; 2.06: 30; 2.07: 25; 2.08: 25 | 1
| 1 | 2 | 2.01: 80.324; 2.02: 20.3; 2.04: < 300; 2.05: < 100; 2.06: 30; 2.07: 25; 2.08: 25 | 4
| 1 | 3 | 2.01: 198.8219; 2.02: 25.4; 2.04: < 300; 2.05: < 100; 2.06: 30; 2.07: 25; 2.08: 25 | 16
| 1 | 4 | 2.01: 39.1148; 2.02: 26.7; 2.04: < 300; 2.05: < 100; 2.06: 30; 2.07: 25; 2.08: 24 | 25.7208
| 1 | 5 | 2.01: 60.3717; 2.02: 20.2; 2.04: < 300; 2.05: < 100; 2.06: 30; 2.07: 25; 2.08: 24 | 10
| 1 | 6 | 2.01: 37.5996; 2.02: 25.5; 2.04: < 300; 2.05: < 100; 2.06: 30; 2.07: 25; 2.08: 24
| Mean |  | 2.01: 79.3789; 2.02: 23.3833; 2.06: 30; 2.07: 25; 2.08: 24.5
| Median |  | 2.01: 60.2064; 2.02: 23.8; 2.06: 30; 2.07: 25; 2.08: 24.5
| Std Dev. |  | 2.01: 60.6126; 2.02: 2.8492; 2.06: 0; 2.07: 0; 2.08: 0.54772
| Rel.Std.Dev. |  | 2.01: 76.36%; 2.02: 12.18%; 2.06: 0%; 2.07: 0%; 2.08: 2.236%
| PDM3 |  | 2.01: 208.6%; 2.02: -9.088%; 2.06: 16.64%; 2.07: -2.803%; 2.08: -4.746%
| Z-Score (Absolute) |  | 2.01: 5.03; 2.02: 0.67; 2.04: 12.23; 2.05: 2.04; 2.06: 0; 2.07: 0.51; 2.08: 0.56
| Analytical results for Pr in OREAS 466 (Certified Value 44.4 ppm) | [19 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  |  | 1
2
| 1 | 1 | 45.6 | 39 | 45.261 | 45.6 | 48.38 | 43.71 | 43 | 42.2 | 41.8 | 47.8 | 44 | 46.4 | 42.3 | 49.7 | 45.8 | 43.6285 |  |  | 1
| 1 | 2 | 45 | 39 | 44.1387 | 41.5 | 48.56 | 43.96 | 41.3 | 43.3 | 45.9 | 46.8 | 45 | 45 | 46 | 48.1 | 45.31 | 47.3414 |  |  | 0
| 1 | 3 | 45.7 | 39 | 45.1274 | 43.3 | 48.71 | 43.34 | 41.5 | 42.6 | 42.3 | 46.7 | 44.4 | 42.6 | 40.9 | 48.6 | 43.84 | 44.7467 |  |  | 16
| 1 | 4 | 44.5 | 42 | 45.0795 | 43 | 47.86 | 42.04 | 43.7 | 41.5 | 43.3 | 47.6 | 45.1 | 45 | 40.4 | 48.5 | 44.74 | 45.1338 |  |  | 44.3622
| 1 | 5 | 44.4 | 39 | 44.6522 | 42.1 | 48.74 | 43.23 | 42.9 | 42.5 | 44.1 | 48.3 | 45 | 43.3 | 44.2 | 48.5 | 43.87 | 44.8779 |  |  | 116
| 1 | 6 | 44.7 | 41 | 45.7252 | 41.9 | 48.96 | 41.74 | 43.9 | 43.8 | 41.5 | 46.7 | 44.5 | 44.7 | 40.4 | 48.9 | 43.24 | 44.2316
| Mean |  | 44.9833 | 39.8333 | 44.9973 | 42.9 | 48.535 | 43.0033 | 42.7167 | 42.65 | 43.15 | 47.3167 | 44.6667 | 44.5 | 42.3667 | 48.7167 | 44.4667 | 44.9933
| Median |  | 44.85 | 39 | 45.1034 | 42.55 | 48.635 | 43.285 | 42.95 | 42.55 | 42.8 | 47.2 | 44.75 | 44.85 | 41.6 | 48.55 | 44.305 | 44.8123
| Std Dev. |  | 0.55648 | 1.3292 | 0.54395 | 1.4873 | 0.38303 | 0.90591 | 1.0926 | 0.81179 | 1.6586 | 0.67946 | 0.43665 | 1.3565 | 2.3001 | 0.54559 | 0.98154 | 1.2689
| Rel.Std.Dev. |  | 1.237% | 3.337% | 1.209% | 3.467% | 0.7892% | 2.107% | 2.558% | 1.903% | 3.844% | 1.436% | 0.9776% | 3.048% | 5.429% | 1.12% | 2.207% | 2.82%
| PDM3 |  | 1.4% | -10.21% | 1.432% | -3.296% | 9.406% | -3.063% | -3.709% | -3.86% | -2.733% | 6.66% | 0.6863% | 0.3106% | -4.498% | 9.816% | 0.2354% | 1.423%
| Z-Score (Absolute) |  | 0.22 | 2.05 | 0.23 | 0.7 | 1.78 | 0.65 | 0.78 | 0.81 | 0.59 | 1.25 | 0.08 | 0.01 | 0.93 | 1.86 | 0.01 | 0.22
| Analytical results for Pr6O11 in OREAS 466 (Certified Value 54 ppm) | [19 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  |  | 0
1
| 1 | 1 | 55.09 | 47.1 | 54.6843 | 55.09 | 58.453 | 52.81 | 51.95 | 50.99 | 50.5 | 57.75 | 53.16 | 56.06 | 51.107 | 60.05 | 55.3356 | 52.7119 |  |  | 1
| 1 | 2 | 54.37 | 47.1 | 53.3283 | 50.14 | 58.67 | 53.112 | 49.9 | 52.32 | 55.46 | 56.54 | 54.37 | 54.369 | 55.577 | 58.11 | 54.7435 | 57.1979 |  |  | 14
| 1 | 3 | 55.21 | 47.1 | 54.5229 | 52.32 | 58.851 | 52.363 | 50.14 | 51.47 | 51.11 | 56.42 | 53.64 | 51.469 | 49.415 | 58.72 | 52.9675 | 54.0629 |  |  | 16
| 1 | 4 | 53.76 | 50.7 | 54.465 | 51.95 | 57.824 | 50.793 | 52.8 | 50.14 | 52.32 | 57.51 | 54.49 | 54.369 | 48.811 | 58.6 | 54.0549 | 54.5307 |  |  | 53.5984
| 1 | 5 | 53.64 | 47.1 | 53.9488 | 50.87 | 58.888 | 52.23 | 51.83 | 51.35 | 53.28 | 58.36 | 54.37 | 52.315 | 53.402 | 58.6 | 53.0037 | 54.2215 |  |  | 117
| 1 | 6 | 54.01 | 49.5 | 55.2452 | 50.62 | 59.153 | 50.43 | 53.04 | 52.92 | 50.14 | 56.42 | 53.76 | 54.007 | 48.811 | 59.08 | 52.2426 | 53.4407
| Mean |  | 54.3467 | 48.1 | 54.3658 | 51.8317 | 58.6398 | 51.9563 | 51.61 | 51.5317 | 52.135 | 57.1667 | 53.965 | 53.7648 | 51.1872 | 58.86 | 53.7246 | 54.3609
| Median |  | 54.19 | 47.1 | 54.494 | 51.41 | 58.7605 | 52.2965 | 51.89 | 51.41 | 51.715 | 57.025 | 54.065 | 54.188 | 50.261 | 58.66 | 53.5293 | 54.1422
| Std Dev. |  | 0.67161 | 1.595 | 0.6572 | 1.7958 | 0.4628 | 1.0943 | 1.3198 | 0.98142 | 2.0052 | 0.8234 | 0.52903 | 1.6389 | 2.779 | 0.66058 | 1.1859 | 1.5331
| Rel.Std.Dev. |  | 1.236% | 3.316% | 1.209% | 3.465% | 0.7892% | 2.106% | 2.557% | 1.904% | 3.846% | 1.44% | 0.9803% | 3.048% | 5.429% | 1.122% | 2.207% | 2.82%
| PDM3 |  | 1.396% | -10.26% | 1.432% | -3.296% | 9.406% | -3.064% | -3.71% | -3.856% | -2.73% | 6.657% | 0.6839% | 0.3104% | -4.499% | 9.817% | 0.2354% | 1.423%
| Z-Score (Absolute) |  | 0.22 | 2.05 | 0.23 | 0.7 | 1.78 | 0.65 | 0.78 | 0.81 | 0.59 | 1.25 | 0.08 | 0.01 | 0.93 | 1.86 | 0.01 | 0.22
| Analytical results for Rb in OREAS 466 (Certified Value 138 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.25 | 2.26 | [5 empty] | ppm
|  |  | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | [5 empty] | 0
0
| 1 | 1 | 129 | 143.7565 | 132 | 142.6 | 122 | 143 | 135 | 141 | 138 | 135 | 151 | 134.6 | 123.126 | [5 empty] | 1
| 1 | 2 | 127 | 140.476 | 131 | 141 | 119 | 141 | 149 | 138 | 138 | 133 | 151 | 132.9 | 132.1727 | [5 empty] | 0
| 1 | 3 | 148 | 141.5883 | 131 | 144.3 | 118 | 146 | 137 | 136 | 139 | 138 | 145 | 136.9 | 126.1607 | [5 empty] | 16
| 1 | 4 | 133 | 142.2861 | 132 | 144.4 | 128 | 142 | 144 | 140 | 140 | 137 | 145 | 133.8 | 126.2009 | [5 empty] | 137.7093
| 1 | 5 | 123 | 138.8674 | 135 | 140.6 | 126 | 152 | 148 | 132 | 135 | 139 | 148 | 135 | 122.9069 | [5 empty] | 118
| 1 | 6 | 131 | 143.6939 | 131 | 138.3 | 124 | 150 | 137 | 140 | 143 | 137 | 139 | 137.8 | 126.0354
| Mean |  | 131.8333 | 141.7781 | 132 | 141.8667 | 122.8333 | 145.6667 | 141.6667 | 137.8333 | 138.8333 | 136.5 | 146.5 | 135.1667 | 126.1004
| Median |  | 130 | 141.9372 | 131.5 | 141.8 | 123 | 144.5 | 140.5 | 139 | 138.5 | 137 | 146.5 | 134.8 | 126.098
| Std Dev. |  | 8.6352 | 1.8998 | 1.5492 | 2.3645 | 3.92 | 4.5019 | 6.121 | 3.3714 | 2.6394 | 2.1679 | 4.5497 | 1.8597 | 3.3448
| Rel.Std.Dev. |  | 6.55% | 1.34% | 1.174% | 1.667% | 3.191% | 3.091% | 4.321% | 2.446% | 1.901% | 1.588% | 3.106% | 1.376% | 2.652%
| PDM3 |  | -4.267% | 2.955% | -4.146% | 3.019% | -10.8% | 5.778% | 2.874% | 0.09006% | 0.8162% | -0.8782% | 6.384% | -1.846% | -8.43%
| Z-Score (Absolute) |  | 1 | 0.66 | 0.98 | 0.67 | 2.51 | 1.31 | 0.64 | 0 | 0.17 | 0.22 | 1.45 | 0.45 | 1.96
| Analytical results for Re in OREAS 466 (Indicative Value 0.004 ppm) | [19 empty] | Indicative
Re | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.17; 2.06: 2.23 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES/MS; 2.06: PF*MS |  |  | 3
3
| 1 | 1 | 2.01: 0.00396; 2.02: < 0.1; 2.06: < 1 |  |  | 1
| 1 | 2 | 2.01: 0.00291; 2.02: < 0.1; 2.06: < 1 |  |  | 1
| 1 | 3 | 2.01: 0.00291; 2.02: < 0.1; 2.06: < 1 |  |  | 16
| 1 | 4 | 2.01: 0.006; 2.02: < 0.1; 2.06: < 1 |  |  | 0.0039483
| 1 | 5 | 2.01: 0.00297; 2.02: < 0.1; 2.06: < 1 |  |  | 11
| 1 | 6 | 2.01: 0.00494; 2.02: < 0.1; 2.06: < 1
| Mean |  | 2.01: 0.0039483
| Median |  | 2.01: 0.003465
| Std Dev. |  | 2.01: 0.0012889
| Rel.Std.Dev. |  | 2.01: 32.64%
| PDM3 |  | 2.01: 8.882e-14%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0; 2.06: 6.59
| Analytical results for S in OREAS 466 (Indicative Value 0.011 wt.%) | [19 empty] | Indicative
S | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.06: 2.05; 2.07: 2.07; 2.08: 2.09; 2.09: 2.17; 2.10: 2.23 |  |  | wt.%
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.06: PF*OES; 2.07: PF*OES/MS; 2.08: PF*OES; 2.09: PF*OES/MS; 2.10: PF*OES |  |  | 3
3
| 1 | 1 | 2.01: < 0.02; 2.02: 0.017471; 2.06: 4.0932e-06; 2.07: < 0.01; 2.08: < 0.01; 2.09: < 0.05; 2.10: < 0.01 |  |  | 1
| 1 | 2 | 2.01: 0.0281; 2.02: 0.017071; 2.06: 4.044e-06; 2.07: < 0.01; 2.08: < 0.01; 2.09: < 0.05; 2.10: < 0.01 |  |  | 6
| 1 | 3 | 2.01: < 0.02; 2.02: 0.017093; 2.06: 8.9656e-06; 2.07: < 0.01; 2.08: < 0.01; 2.09: < 0.05; 2.10: < 0.01 |  |  | 16
| 1 | 4 | 2.01: 0.0498; 2.02: 0.016856; 2.06: 3.5026e-06; 2.07: < 0.01; 2.08: 0.01; 2.09: < 0.05; 2.10: < 0.01 |  |  | 0.010505
| 1 | 5 | 2.01: < 0.02; 2.02: 0.017004; 2.06: 3.3549e-06; 2.07: < 0.01; 2.08: 0.01; 2.09: < 0.05; 2.10: < 0.01 |  |  | 12
| 1 | 6 | 2.01: 0.0333; 2.02: 0.016914; 2.06: 2.8053e-06; 2.07: < 0.01; 2.08: < 0.01; 2.09: < 0.05; 2.10: < 0.01
| Mean |  | 2.01: 0.037067; 2.02: 0.017068; 2.06: 4.4609e-06; 2.08: 0.01
| Median |  | 2.01: 0.0333; 2.02: 0.017038; 2.06: 3.7733e-06; 2.08: 0.01
| Std Dev. |  | 2.01: 0.01133; 2.02: 0.00021716; 2.06: 2.2576e-06; 2.08: 0
| Rel.Std.Dev. |  | 2.01: 30.57%; 2.02: 1.272%; 2.06: 50.61%; 2.08: 0%
| PDM3 |  | 2.01: 252.9%; 2.02: 62.48%; 2.06: -99.96%; 2.08: -4.804%
| Z-Score (Absolute) |  | 2.01: 1.71; 2.02: 1.05; 2.06: 0.67; 2.07: 0.17; 2.08: 0; 2.09: 1.86; 2.10: 0.17
| Analytical results for Sb in OREAS 466 (Indicative Value 5.78 ppm) | [19 empty] | Indicative
Sb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab |  |  | 1
| No | No | 2.01: 2.05; 2.02: 2.07; 2.06: 2.17; 2.07: 2.23; 2.08: 2.25 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS |  |  | 2
2
| 1 | 1 | 2.01: 2.0601; 2.02: 6; 2.06: 6.1; 2.07: 5.9; 2.08: 5.9 |  |  | 1
| 1 | 2 | 2.01: 1.4789; 2.02: 6; 2.06: 6; 2.07: 5.6; 2.08: 5.6 |  |  | 7
| 1 | 3 | 2.01: 1.7713; 2.02: 6; 2.06: 6; 2.07: 5.2; 2.08: 5.8 |  |  | 16
| 1 | 4 | 2.01: 2.7267; 2.02: 6; 2.06: 5.9; 2.07: 5.9; 2.08: 5.5 |  |  | 5.7792
| 1 | 5 | 2.01: 2.1374; 2.02: 5; 2.06: 6; 2.07: 5.9; 2.08: 5.7 |  |  | 13
| 1 | 6 | 2.01: 1.6824; 2.02: 7; 2.06: 6.1; 2.07: 5.3; 2.08: 5.5
| Mean |  | 2.01: 1.9761; 2.02: 6; 2.06: 6.0167; 2.07: 5.6333; 2.08: 5.6667
| Median |  | 2.01: 1.9157; 2.02: 6; 2.06: 6; 2.07: 5.75; 2.08: 5.65
| Std Dev. |  | 2.01: 0.44088; 2.02: 0.63246; 2.06: 0.075277; 2.07: 0.32042; 2.08: 0.1633
| Rel.Std.Dev. |  | 2.01: 22.31%; 2.02: 10.54%; 2.06: 1.251%; 2.07: 5.688%; 2.08: 2.882%
| PDM3 |  | 2.01: -65.81%; 2.02: 3.821%; 2.06: 4.11%; 2.07: -2.523%; 2.08: -1.947%
| Z-Score (Absolute) |  | 2.01: 7.47; 2.02: 0.67; 2.06: 0.71; 2.07: 0.07; 2.08: 0
| Analytical results for Sc in OREAS 466 (Certified Value 19.2 ppm) | [19 empty] | Certified
Sc | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.08; 2.06: 2.17; 2.07: 2.23; 2.08: 2.24; 2.09: 2.25; 2.10: 2.26 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.06: PF*OES/MS; 2.07: PF*OES; 2.08: PF*MS; 2.09: PF*MS; 2.10: BF*MS |  |  | 1
1
| 1 | 1 | 2.01: 22; 2.02: 24.03; 2.06: < 20; 2.07: 21; 2.08: 14.8; 2.09: 17; 2.10: 19.3472 |  |  | 1
| 1 | 2 | 2.01: 22; 2.02: 23.1; 2.06: < 20; 2.07: 20; 2.08: 14.4; 2.09: 17; 2.10: 18.9747 |  |  | 0
| 1 | 3 | 2.01: 22; 2.02: 23.83; 2.06: < 20; 2.07: 17; 2.08: 14.6; 2.09: 17; 2.10: 17.8485 |  |  | 16
| 1 | 4 | 2.01: 21; 2.02: 23.34; 2.06: < 20; 2.07: 18; 2.08: 15; 2.09: 17; 2.10: 18.1771 |  |  | 19.1837
| 1 | 5 | 2.01: 22; 2.02: 25; 2.06: < 20; 2.07: 21; 2.08: 14.6; 2.09: 17; 2.10: 17.7377 |  |  | 119
| 1 | 6 | 2.01: 22; 2.02: 24.8; 2.06: < 20; 2.07: 19; 2.08: 14.2; 2.09: 17; 2.10: 17.828
| Mean |  | 2.01: 21.8333; 2.02: 24.0167; 2.07: 19.3333; 2.08: 14.6; 2.09: 17; 2.10: 18.3189
| Median |  | 2.01: 22; 2.02: 23.93; 2.07: 19.5; 2.08: 14.6; 2.09: 17; 2.10: 18.0128
| Std Dev. |  | 2.01: 0.40825; 2.02: 0.76338; 2.07: 1.633; 2.08: 0.28284; 2.09: 0; 2.10: 0.67935
| Rel.Std.Dev. |  | 2.01: 1.87%; 2.02: 3.179%; 2.07: 8.447%; 2.08: 1.937%; 2.09: 0%; 2.10: 3.708%
| PDM3 |  | 2.01: 13.81%; 2.02: 25.19%; 2.07: 0.78%; 2.08: -23.89%; 2.09: -11.38%; 2.10: -4.508%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 1.09; 2.06: 1.6; 2.07: 0.19; 2.08: 0.71; 2.09: 0.25; 2.10: 0
| Analytical results for Se in OREAS 466 (Indicative Value 13.7 ppm) | [19 empty] | Indicative
Se | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab |  |  | 1
| No | No | 2.01: 2.05; 2.02: 2.07; 2.06: 2.17; 2.07: 2.23 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS |  |  | 1
1
| 1 | 1 | 2.02: 21; 2.06: < 20; 2.07: < 50 |  |  | 1
| 1 | 2 | 2.02: 16; 2.06: < 20; 2.07: < 50 |  |  | 2
| 1 | 3 | 2.01: 10.1915; 2.02: 12; 2.06: < 20; 2.07: < 50 |  |  | 16
| 1 | 4 | 2.02: 17; 2.06: < 20; 2.07: < 50 |  |  | 13.7217
| 1 | 5 | 2.02: 18; 2.06: < 20; 2.07: < 50 |  |  | 10
| 1 | 6 | 2.01: 11.6951; 2.02: 15; 2.06: < 20; 2.07: < 50
| Mean |  | 2.01: 10.9433; 2.02: 16.5
| Median |  | 2.01: 10.9433; 2.02: 16.5
| Std Dev. |  | 2.01: 1.0632; 2.02: 3.0166
| Rel.Std.Dev. |  | 2.01: 9.716%; 2.02: 18.28%
| PDM3 |  | 2.01: -20.25%; 2.02: 20.25%
| Z-Score (Absolute) |  | 2.01: 0.58; 2.02: 0.58; 2.06: 0.77; 2.07: 2.34
| Analytical results for Si in OREAS 466 (Certified Value 30.7 wt.%) | [19 empty] | Certified
Si | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.06: 2.05; 2.07: 2.07; 2.08: 2.15; 2.09: 2.17; 2.10: 2.23; 2.11: 2.26 |  |  | wt.%
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.06: PF*OES; 2.07: PF*OES/MS; 2.08: PF*OES; 2.09: PF*OES/MS; 2.10: PF*OES; 2.11: BF*OES |  |  | 2
3
| 1 | 1 | 2.01: 26.43; 2.02: 31.5514; 2.06: 31.5634; 2.07: > 30; 2.08: 28.8; 2.09: 28.1; 2.10: 32.1; 2.11: 31.5418 |  |  | 1
| 1 | 2 | 2.01: 26.2834; 2.02: 31.2921; 2.06: 32.2537; 2.07: > 30; 2.08: 29.6; 2.09: 28.3; 2.10: 32.2; 2.11: 31.9605 |  |  | 0
| 1 | 3 | 2.01: 26.7696; 2.02: 31.7531; 2.06: 31.3876; 2.07: > 30; 2.08: 28.1; 2.09: 27.9; 2.10: 31.8; 2.11: 31.7837 |  |  | 16
| 1 | 4 | 2.01: 25.2342; 2.02: 31.9468; 2.06: 32.3766; 2.07: > 30; 2.08: 28.5; 2.09: 28.3; 2.10: 31.1; 2.11: 32.1024 |  |  | 30.6954
| 1 | 5 | 2.01: 26.6374; 2.02: 31.1832; 2.06: 31.7767; 2.07: > 30; 2.08: 29.2; 2.09: 28.2; 2.10: 31.9; 2.11: 31.9907 |  |  | 120
| 1 | 6 | 2.01: 26.9841; 2.02: 31.6034; 2.06: 31.7388; 2.07: > 30; 2.08: 29.2; 2.09: 28.1; 2.10: 32; 2.11: 31.8293
| Mean |  | 2.01: 26.3898; 2.02: 31.555; 2.06: 31.8495; 2.08: 28.9; 2.09: 28.15; 2.10: 31.85; 2.11: 31.8681
| Median |  | 2.01: 26.5337; 2.02: 31.5774; 2.06: 31.7578; 2.08: 29; 2.09: 28.15; 2.10: 31.95; 2.11: 31.8949
| Std Dev. |  | 2.01: 0.61759; 2.02: 0.28371; 2.06: 0.38832; 2.08: 0.54406; 2.09: 0.15166; 2.10: 0.3937; 2.11: 0.19677
| Rel.Std.Dev. |  | 2.01: 2.34%; 2.02: 0.8991%; 2.06: 1.219%; 2.08: 1.883%; 2.09: 0.5387%; 2.10: 1.236%; 2.11: 0.6175%
| PDM3 |  | 2.01: -14.03%; 2.02: 2.8%; 2.06: 3.76%; 2.08: -5.849%; 2.09: -8.293%; 2.10: 3.761%; 2.11: 3.82%
| Z-Score (Absolute) |  | 2.01: 11.13; 2.02: 0; 2.06: 0.63; 2.07: NA; 2.08: 5.72; 2.09: 7.33; 2.10: 0.64; 2.11: 0.67
| Analytical results for SiO2 in OREAS 466 (Certified Value 65.66 wt.%) | [19 empty] | Certified
SiO2 | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.06: 2.05; 2.07: 2.07; 2.08: 2.15; 2.09: 2.17; 2.10: 2.23; 2.11: 2.26 |  |  | wt.%
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.06: PF*OES; 2.07: PF*OES/MS; 2.08: PF*OES; 2.09: PF*OES/MS; 2.10: PF*OES; 2.11: BF*OES |  |  | 2
3
| 1 | 1 | 2.01: 56.5391; 2.02: 67.4947; 2.06: 67.5204; 2.07: > 64; 2.08: 61.61; 2.09: 60.11; 2.10: 68.67; 2.11: 67.4742 |  |  | 1
| 1 | 2 | 2.01: 56.2255; 2.02: 66.9401; 2.06: 68.9971; 2.07: > 64; 2.08: 63.32; 2.09: 60.54; 2.10: 68.88; 2.11: 68.3698 |  |  | 0
| 1 | 3 | 2.01: 57.2655; 2.02: 67.9262; 2.06: 67.1443; 2.07: > 64; 2.08: 60.11; 2.09: 59.68; 2.10: 68.03; 2.11: 67.9916 |  |  | 16
| 1 | 4 | 2.01: 53.981; 2.02: 68.3405; 2.06: 69.26; 2.07: > 64; 2.08: 60.97; 2.09: 60.54; 2.10: 66.53; 2.11: 68.6735 |  |  | 65.6636
| 1 | 5 | 2.01: 56.9827; 2.02: 66.707; 2.06: 67.9768; 2.07: > 64; 2.08: 62.46; 2.09: 60.33; 2.10: 68.24; 2.11: 68.4346 |  |  | 121
| 1 | 6 | 2.01: 57.7244; 2.02: 67.606; 2.06: 67.8957; 2.07: > 64; 2.08: 62.46; 2.09: 60.11; 2.10: 68.45; 2.11: 68.0892
| Mean |  | 2.01: 56.453; 2.02: 67.5024; 2.06: 68.1324; 2.08: 61.8217; 2.09: 60.2183; 2.10: 68.1333; 2.11: 68.1721
| Median |  | 2.01: 56.7609; 2.02: 67.5503; 2.06: 67.9363; 2.08: 62.035; 2.09: 60.22; 2.10: 68.345; 2.11: 68.2295
| Std Dev. |  | 2.01: 1.3212; 2.02: 0.60691; 2.06: 0.8307; 2.08: 1.1627; 2.09: 0.3264; 2.10: 0.84125; 2.11: 0.42093
| Rel.Std.Dev. |  | 2.01: 2.34%; 2.02: 0.8991%; 2.06: 1.219%; 2.08: 1.881%; 2.09: 0.542%; 2.10: 1.235%; 2.11: 0.6175%
| PDM3 |  | 2.01: -14.03%; 2.02: 2.8%; 2.06: 3.76%; 2.08: -5.851%; 2.09: -8.293%; 2.10: 3.761%; 2.11: 3.82%
| Z-Score (Absolute) |  | 2.01: 11.13; 2.02: 0; 2.06: 0.63; 2.07: NA; 2.08: 5.72; 2.09: 7.33; 2.10: 0.64; 2.11: 0.67
| Analytical results for Sm in OREAS 466 (Certified Value 21 ppm) | [19 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.23 | 2.24 | 2.25 | 2.26 |  |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  |  | 1
2
| 1 | 1 | 21.5 | 19 | 21.9505 | 19.7 | 24.54 | 20.7 | 23 | 19.9 | 20.9 | 23.8 | 21.8 | 20.5 | 21.4 | 21.1 | 20.5887 |  |  |  | 1
| 1 | 2 | 21 | 19 | 20.5788 | 19.9 | 24.89 | 20.97 | 20.2 | 20 | 22.3 | 22.6 | 21.3 | 19.5 | 21.1 | 21.6 | 23.2964 |  |  |  | 0
| 1 | 3 | 21.6 | 19 | 22.751 | 19.3 | 24.92 | 20.25 | 19.3 | 19.8 | 21.1 | 23.4 | 22 | 20.8 | 21.2 | 21.1 | 21.6058 |  |  |  | 16
| 1 | 4 | 21.7 | 19 | 22.1091 | 18.9 | 24.42 | 20.52 | 24 | 19.4 | 21 | 23 | 22 | 19.1 | 20.7 | 20.9 | 20.8172 |  |  |  | 20.9838
| 1 | 5 | 20.6 | 19 | 21.3551 | 20.3 | 24.61 | 20.35 | 23.2 | 20.2 | 22.8 | 23 | 22.6 | 19.3 | 20.2 | 21.6 | 20.8001 |  |  |  | 122
| 1 | 6 | 21.1 | 20 | 22.1441 | 18.8 | 23.9 | 20.24 | 20.8 | 21.6 | 20.1 | 23.5 | 22.2 | 19.7 | 20.2 | 21.2 | 20.2155
| Mean |  | 21.25 | 19.1667 | 21.8148 | 19.4833 | 24.5467 | 20.505 | 21.75 | 20.15 | 21.3667 | 23.2167 | 21.9833 | 19.8167 | 20.8 | 21.25 | 21.2206
| Median |  | 21.3 | 19 | 22.0298 | 19.5 | 24.575 | 20.435 | 21.9 | 19.95 | 21.05 | 23.2 | 22 | 19.6 | 20.9 | 21.15 | 20.8087
| Std Dev. |  | 0.42308 | 0.40825 | 0.75219 | 0.58793 | 0.37297 | 0.28766 | 1.8992 | 0.75829 | 0.99532 | 0.43089 | 0.43089 | 0.6824 | 0.51769 | 0.2881 | 1.1141
| Rel.Std.Dev. |  | 1.991% | 2.13% | 3.448% | 3.018% | 1.519% | 1.403% | 8.732% | 3.763% | 4.658% | 1.856% | 1.96% | 3.444% | 2.489% | 1.356% | 5.25%
| PDM3 |  | 1.268% | -8.66% | 3.96% | -7.151% | 16.98% | -2.282% | 3.651% | -3.974% | 1.824% | 10.64% | 4.763% | -5.562% | -0.8761% | 1.268% | 1.128%
| Z-Score (Absolute) |  | 0 | 1.92 | 0.52 | 1.62 | 3.03 | 0.69 | 0.46 | 1.01 | 0.11 | 1.81 | 0.67 | 1.32 | 0.41 | 0 | 0.03
| Analytical results for Sm2O3 in OREAS 466 (Certified Value 24.3 ppm) | [19 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.23 | 2.24 | 2.25 | 2.26 |  |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  |  | 1
2
| 1 | 1 | 24.93 | 22 | 25.4538 | 22.84 | 28.457 | 24.004 | 26.67 | 23.08 | 24.24 | 27.6 | 25.28 | 23.772 | 24.82 | 24.468 | 23.8747 |  |  |  | 1
| 1 | 2 | 24.35 | 22 | 23.8632 | 23.08 | 28.862 | 24.317 | 23.42 | 23.19 | 25.86 | 26.21 | 24.7 | 22.612 | 24.47 | 25.047 | 27.0145 |  |  |  | 0
| 1 | 3 | 25.05 | 22 | 26.3821 | 22.38 | 28.897 | 23.482 | 22.38 | 22.96 | 24.47 | 27.13 | 25.51 | 24.12 | 24.58 | 24.468 | 25.054 |  |  |  | 16
| 1 | 4 | 25.16 | 22 | 25.6377 | 21.92 | 28.317 | 23.795 | 27.83 | 22.5 | 24.35 | 26.67 | 25.51 | 22.148 | 24 | 24.236 | 24.1397 |  |  |  | 24.3329
| 1 | 5 | 23.89 | 22 | 24.7633 | 23.54 | 28.538 | 23.598 | 26.9 | 23.42 | 26.44 | 26.67 | 26.21 | 22.38 | 23.42 | 25.047 | 24.1198 |  |  |  | 123
| 1 | 6 | 24.47 | 23.2 | 25.6783 | 21.8 | 27.714 | 23.47 | 24.12 | 25.05 | 23.31 | 27.25 | 25.74 | 22.844 | 23.42 | 24.584 | 23.4419
| Mean |  | 24.6417 | 22.2 | 25.2964 | 22.5933 | 28.4642 | 23.7777 | 25.22 | 23.3667 | 24.7783 | 26.9217 | 25.4917 | 22.9793 | 24.1183 | 24.6417 | 24.6074
| Median |  | 24.7 | 22 | 25.5458 | 22.61 | 28.4975 | 23.6965 | 25.395 | 23.135 | 24.41 | 26.9 | 25.51 | 22.728 | 24.235 | 24.526 | 24.1297
| Std Dev. |  | 0.48951 | 0.4899 | 0.87224 | 0.68149 | 0.43256 | 0.3337 | 2.2024 | 0.87936 | 1.1537 | 0.4989 | 0.49998 | 0.79158 | 0.60307 | 0.33371 | 1.2919
| Rel.Std.Dev. |  | 1.986% | 2.207% | 3.448% | 3.016% | 1.52% | 1.403% | 8.733% | 3.763% | 4.656% | 1.853% | 1.961% | 3.445% | 2.5% | 1.354% | 5.25%
| PDM3 |  | 1.269% | -8.765% | 3.96% | -7.149% | 16.98% | -2.282% | 3.646% | -3.971% | 1.831% | 10.64% | 4.762% | -5.563% | -0.8816% | 1.269% | 1.128%
| Z-Score (Absolute) |  | 0 | 1.92 | 0.52 | 1.62 | 3.03 | 0.69 | 0.46 | 1.01 | 0.11 | 1.81 | 0.67 | 1.32 | 0.41 | 0 | 0.03
| Analytical results for Sn in OREAS 466 (Certified Value 20.5 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.23 | 2.25 | 2.26 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | PF*MS | BF*MS | BF*MS |  |  |  |  | 1
2
| 1 | 1 | 20 | 35.3549 | 20.9932 | 18.5 | 20 | 20 | 18 | 21 | 21 | 20 | < 50 | 20 | 20.8 | 22.9674 |  |  |  |  | 1
| 1 | 2 | 19 | 48.4597 | 19.7135 | 19.2 | 21 | 20 | 18 | 22 | 22 | 22 | < 50 | 21 | 20.9 | 24.0986 |  |  |  |  | 0
| 1 | 3 | 19 | 54.3675 | 20.1486 | 19.7 | 20 | 20 | 19 | 21 | 22 | 22 | 80 | 21 | 21.3 | 21.0323 |  |  |  |  | 16
| 1 | 4 | 22 | 53.7272 | 20.542 | 19.5 | 20 | 23 | 18 | 22 | 21 | 21 | < 50 | 21 | 21.1 | 20.8724 |  |  |  |  | 20.4984
| 1 | 5 | 18 | 34.9465 | 20.6677 | 19.8 | 20 | 24 | 19 | 21 | 22 | 17 | < 50 | 21 | 20.8 | 20.1343 |  |  |  |  | 124
| 1 | 6 | 30 | 33.5458 | 20.5242 | 19.6 | 20 | 21 | 20 | 20 | 21 | 20 | < 50 | 20 | 20.9 | 21.4891
| Mean |  | 21.3333 | 43.4003 | 20.4316 | 19.3833 | 20.1667 | 21.3333 | 18.6667 | 21.1667 | 21.5 | 20.3333 | 80 | 20.6667 | 20.9667 | 21.7657
| Median |  | 19.5 | 41.9073 | 20.5331 | 19.55 | 20 | 20.5 | 18.5 | 21 | 21.5 | 20.5 | 80 | 21 | 20.9 | 21.2607
| Std Dev. |  | 4.4572 | 9.8573 | 0.44444 | 0.47924 | 0.40825 | 1.7512 | 0.8165 | 0.75277 | 0.54772 | 1.8619 |  | 0.5164 | 0.19664 | 1.4806
| Rel.Std.Dev. |  | 20.89% | 22.71% | 2.175% | 2.472% | 2.024% | 8.209% | 4.374% | 3.556% | 2.548% | 9.157% |  | 2.499% | 0.9379% | 6.802%
| PDM3 |  | 4.073% | 111.7% | -0.326% | -5.44% | -1.618% | 4.073% | -8.936% | 3.26% | 4.886% | -0.8052% | 290.3% | 0.821% | 2.285% | 6.182%
| Z-Score (Absolute) |  | 0.27 | 22.58 | 0.64 | 1.7 | 0.91 | 0.27 | 2.43 | 0.1 | 0.44 | 0.74 | 13.24 | 0.4 | 0.1 | 0.71
| Analytical results for Sr in OREAS 466 (Certified Value 110 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES | [7 empty] | 0
0
| 1 | 1 | 132 | 109.2225 | 123.2525 | 115.4806 | 103 | 113.4 | 110 | 120 | 117 | 102 | 108.365 | [7 empty] | 1
| 1 | 2 | 129 | 119.0075 | 133.3272 | 116.1828 | 101 | 113.7 | 130 | 120 | 110 | 103 | 111.395 | [7 empty] | 0
| 1 | 3 | 136 | 114.039 | 132.0332 | 115.7408 | 100 | 115.8 | 110 | 110 | 117 | 110 | 109.13 | [7 empty] | 16
| 1 | 4 | 198 | 109.5835 | 127.312 | 117.5332 | 103 | 114.7 | 110 | 110 | 115 | 104 | 111.545 | [7 empty] | 109.8755
| 1 | 5 | 132 | 112.614 | 127.0632 | 112.5133 | 100 | 113.5 | 110 | 110 | 111 | 98 | 110.07 | [7 empty] | 125
| 1 | 6 | 132 | 110.619 | 126.6433 | 114.8962 | 100 | 108.8 | 120 | 110 | 106 | 100 | 111.515
| Mean |  | 143.1667 | 112.5143 | 128.2719 | 115.3912 | 101.1667 | 113.3167 | 115 | 113.3333 | 112.6667 | 102.8333 | 110.3367
| Median |  | 132 | 111.6165 | 127.1876 | 115.6107 | 100.5 | 113.6 | 110 | 110 | 113 | 102.5 | 110.7325
| Std Dev. |  | 26.9549 | 3.6746 | 3.7399 | 1.6657 | 1.472 | 2.3945 | 8.3666 | 5.164 | 4.4121 | 4.1191 | 1.3699
| Rel.Std.Dev. |  | 18.83% | 3.266% | 2.916% | 1.443% | 1.455% | 2.113% | 7.275% | 4.556% | 3.916% | 4.006% | 1.242%
| PDM3 |  | 30.3% | 2.402% | 16.74% | 5.02% | -7.926% | 3.132% | 4.664% | 3.147% | 2.54% | -6.409% | 0.4197%
| Z-Score (Absolute) |  | 7.55 | 0.04 | 3.86 | 0.67 | 2.85 | 0.16 | NA | NA | 0 | 2.43 | 0.58
Results from laboratories 15 and 17 were removed due to their 10 ppm reading resolution.
| Analytical results for Ta in OREAS 466 (Certified Value 2.94 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  |  | 2
3
| 1 | 1 | 3 | 2.5157 | 3.0371 | 2.7 | 3 | 2.3 | 2.9 | 2.8 | 3 | 2.7 | 3.7 | < 10 | 5.8 | 3.12 | 2.7349 |  |  |  | 1
| 1 | 2 | 3.2 | 2.4858 | 2.9439 | 2.9 | 3.4 | 2.2 | 3 | 3 | 3.1 | 3.1 | 3.2 | < 10 | 5.4 | 3.69 | 2.8992 |  |  |  | 0
| 1 | 3 | 3 | 2.7179 | 3.1972 | 2.7 | 2.9 | 2.4 | 3.1 | 2.8 | 3 | 2.9 | 3.9 | < 10 | 5.5 | 3.14 | 2.9373 |  |  |  | 16
| 1 | 4 | 2.7 | 2.541 | 3.0825 | 2.6 | 2.7 | 2.4 | 3.1 | 2.9 | 3 | 2.8 | 3.8 | < 10 | 5.6 | 3 | 2.8632 |  |  |  | 2.9421
| 1 | 5 | 3 | 2.3766 | 3.0521 | 2.6 | 3.2 | 2.4 | 3 | 2.9 | 2.8 | 3 | 3.3 | < 10 | 5.6 | 3.29 | 2.7096 |  |  |  | 126
| 1 | 6 | 3 | 2.7197 | 2.8827 | 2.7 | 2.6 | 2.3 | 3 | 2.9 | 2.9 | 2.9 | 3 | < 10 | 5.7 | 4.39 | 2.8598
| Mean |  | 2.9833 | 2.5594 | 3.0326 | 2.7 | 2.9667 | 2.3333 | 3.0167 | 2.8833 | 2.9667 | 2.9 | 3.4833 |  | 5.6 | 3.4383 | 2.834
| Median |  | 3 | 2.5283 | 3.0446 | 2.7 | 2.95 | 2.35 | 3 | 2.9 | 3 | 2.9 | 3.5 |  | 5.6 | 3.215 | 2.8615
| Std Dev. |  | 0.16021 | 0.13555 | 0.10982 | 0.10954 | 0.30111 | 0.08165 | 0.075277 | 0.075277 | 0.10328 | 0.14142 | 0.3656 |  | 0.14142 | 0.52412 | 0.091341
| Rel.Std.Dev. |  | 5.37% | 5.296% | 3.621% | 4.057% | 10.15% | 3.499% | 2.495% | 2.611% | 3.481% | 4.877% | 10.5% |  | 2.525% | 15.24% | 3.223%
| PDM3 |  | 1.402% | -13.01% | 3.077% | -8.228% | 0.8359% | -20.69% | 2.535% | -1.997% | 0.8359% | -1.43% | 18.4% |  | 90.34% | 16.87% | -3.674%
| Z-Score (Absolute) |  | 0.08 | 2.07 | 0.34 | 1.36 | 0 | 3.22 | 0.25 | 0.42 | 0 | 0.34 | 2.63 | 10.33 | 13.38 | 2.4 | 0.67
| Analytical results for Tb in OREAS 466 (Certified Value 1.48 ppm) | [19 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 1.5 | < 2 | 1.4535 | 1.6 | 1.31 | 1.45 | 1.52 | 1.49 | 1.44 | 1.44 | 1.35 | 1.58 | 1.6 | 52.3 | 1.37 | 1.4769 |  |  | 1
| 1 | 2 | 1.6 | < 2 | 1.3969 | 1.5 | 1.33 | 1.49 | 1.32 | 1.49 | 1.49 | 1.34 | 1.49 | 1.55 | 1.8 | 51.7 | 1.42 | 1.5316 |  |  | 0
| 1 | 3 | 1.6 | < 2 | 1.3479 | 1.6 | 1.36 | 1.43 | 1.38 | 1.45 | 1.42 | 1.59 | 1.54 | 1.55 | 1.6 | 52.4 | 1.42 | 1.5258 |  |  | 16
| 1 | 4 | 1.6 | < 2 | 1.4638 | 1.6 | 1.37 | 1.42 | 1.52 | 1.42 | 1.54 | 1.46 | 1.56 | 1.62 | 1.7 | 51.6 | 1.44 | 1.4991 |  |  | 1.4826
| 1 | 5 | 1.5 | < 2 | 1.4659 | 1.7 | 1.31 | 1.46 | 1.38 | 1.43 | 1.6 | 1.46 | 1.45 | 1.5 | 1.8 | 50.9 | 1.36 | 1.4226 |  |  | 127
| 1 | 6 | 1.5 | < 2 | 1.3957 | 1.6 | 1.32 | 1.43 | 1.5 | 1.55 | 1.44 | 1.51 | 1.57 | 1.6 | 1.5 | 51.8 | 1.4 | 1.4263
| Mean |  | 1.55 |  | 1.4206 | 1.6 | 1.3333 | 1.4467 | 1.4367 | 1.4717 | 1.4883 | 1.4667 | 1.4933 | 1.5667 | 1.6667 | 51.7833 | 1.4017 | 1.4804
| Median |  | 1.55 |  | 1.4252 | 1.6 | 1.325 | 1.44 | 1.44 | 1.47 | 1.465 | 1.46 | 1.515 | 1.565 | 1.65 | 51.75 | 1.41 | 1.488
| Std Dev. |  | 0.054772 |  | 0.047899 | 0.063246 | 0.02582 | 0.02582 | 0.087101 | 0.048339 | 0.069976 | 0.082381 | 0.083586 | 0.04274 | 0.12111 | 0.54191 | 0.031252 | 0.047559
| Rel.Std.Dev. |  | 3.534% |  | 3.372% | 3.953% | 1.936% | 1.785% | 6.063% | 3.285% | 4.702% | 5.617% | 5.597% | 2.728% | 7.266% | 1.046% | 2.23% | 3.213%
| PDM3 |  | 4.548% |  | -4.179% | 7.921% | -10.07% | -2.422% | -3.096% | -0.7354% | 0.3888% | -1.073% | 0.726% | 5.672% | 12.42% | 3393% | -5.457% | -0.1485%
| Z-Score (Absolute) |  | 0.77 | 4.96 | 0.58 | 1.29 | 1.49 | 0.31 | 0.41 | 0.05 | 0.13 | 0.1 | 0.18 | 0.94 | 1.99 | 524.34 | 0.77 | 0.05
| Analytical results for Tb4O7 in OREAS 466 (Certified Value 1.74 ppm) | [19 empty] | Certified
Tb4O7 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 1.76 | < 2.4 | 1.7096 | 1.88 | 1.541 | 1.705 | 1.788 | 1.753 | 1.694 | 1.694 | 1.588 | 1.858 | 1.882 | 61.52 | 1.611 | 1.7371 |  |  | 1
| 1 | 2 | 1.88 | < 2.4 | 1.6431 | 1.76 | 1.564 | 1.753 | 1.553 | 1.753 | 1.753 | 1.576 | 1.753 | 1.823 | 2.117 | 60.81 | 1.67 | 1.8015 |  |  | 0
| 1 | 3 | 1.88 | < 2.4 | 1.5853 | 1.88 | 1.6 | 1.682 | 1.623 | 1.705 | 1.67 | 1.87 | 1.811 | 1.823 | 1.882 | 61.63 | 1.67 | 1.7946 |  |  | 16
| 1 | 4 | 1.88 | < 2.4 | 1.7217 | 1.88 | 1.611 | 1.67 | 1.788 | 1.67 | 1.811 | 1.717 | 1.835 | 1.905 | 2 | 60.69 | 1.694 | 1.7632 |  |  | 1.7438
| 1 | 5 | 1.76 | < 2.4 | 1.7242 | 2 | 1.541 | 1.717 | 1.623 | 1.682 | 1.882 | 1.717 | 1.705 | 1.764 | 2.117 | 59.87 | 1.6 | 1.6733 |  |  | 128
| 1 | 6 | 1.76 | < 2.4 | 1.6416 | 1.88 | 1.553 | 1.682 | 1.764 | 1.823 | 1.694 | 1.776 | 1.847 | 1.882 | 1.764 | 60.93 | 1.647 | 1.6776
| Mean |  | 1.82 |  | 1.6709 | 1.88 | 1.5683 | 1.7015 | 1.6898 | 1.731 | 1.7507 | 1.725 | 1.7565 | 1.8425 | 1.9603 | 60.9083 | 1.6487 | 1.7412
| Median |  | 1.82 |  | 1.6763 | 1.88 | 1.5585 | 1.6935 | 1.6935 | 1.729 | 1.7235 | 1.717 | 1.782 | 1.8405 | 1.941 | 60.87 | 1.6585 | 1.7501
| Std Dev. |  | 0.065727 |  | 0.056339 | 0.075895 | 0.030237 | 0.03052 | 0.1024 | 0.056974 | 0.082248 | 0.096867 | 0.098366 | 0.050258 | 0.14246 | 0.63732 | 0.036757 | 0.055938
| Rel.Std.Dev. |  | 3.611% |  | 3.372% | 4.037% | 1.928% | 1.794% | 6.06% | 3.291% | 4.698% | 5.615% | 5.6% | 2.728% | 7.267% | 1.046% | 2.229% | 3.213%
| PDM3 |  | 4.37% |  | -4.179% | 7.811% | -10.06% | -2.426% | -3.095% | -0.7339% | 0.3939% | -1.078% | 0.7284% | 5.66% | 12.42% | 3393% | -5.455% | -0.1485%
| Z-Score (Absolute) |  | 0.77 | 4.96 | 0.58 | 1.29 | 1.49 | 0.31 | 0.41 | 0.05 | 0.13 | 0.1 | 0.18 | 0.94 | 1.99 | 524.34 | 0.77 | 0.05
| Analytical results for Te in OREAS 466 (Indicative Value < 1 ppm) | [19 empty] | Indicative
Te | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.06: Lab |  |  | 1
| No | No | 2.01: 2.07; 2.02: 2.17; 2.06: 2.23 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES/MS; 2.06: PF*MS |  |  | 2
2
| 1 | 1 | 2.01: < 6; 2.02: < 1; 2.06: < 1 |  |  | 1
| 1 | 2 | 2.01: < 6; 2.02: < 1; 2.06: < 1 |  |  | 5
| 1 | 3 | 2.01: < 6; 2.02: < 1; 2.06: < 1 |  |  | 16
| 1 | 4 | 2.01: < 6; 2.02: < 1; 2.06: < 1 |  |  | < 1
| 1 | 5 | 2.01: < 6; 2.02: < 1; 2.06: < 1 |  |  | 11
| 1 | 6 | 2.01: < 6; 2.02: < 1; 2.06: < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.06
| Analytical results for Th in OREAS 466 (Certified Value 31.9 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | 1
2
| 1 | 1 | 32.3 | 33 | 38.9539 | 38.52 | 31.5341 | 32.4 | 30.24 | 30.5 | 33 | 30.3 | 32.1 | 31.4 | 31.1 | 35.5 | 31 | 35.3 | 31.3 | 31.2017 | 1
| 1 | 2 | 31.4 | 33 | 38.6502 | 37.51 | 30.0378 | 32.4 | 30.01 | 31.7 | 29.9 | 30.6 | 34.9 | 31.4 | 31.7 | 36.1 | 32.2 | 35.1 | 30.7 | 33.797 | 0
| 1 | 3 | 31.2 | 33 | 38.6998 | 37.92 | 32.3693 | 32.6 | 30.04 | 30.2 | 31.8 | 31.3 | 30.7 | 32.5 | 31.1 | 35 | 30.6 | 35.1 | 30.4 | 32.2704 | 16
| 1 | 4 | 31.4 | 32 | 38.1339 | 38.9 | 30.4496 | 32.5 | 30.26 | 29.8 | 32.9 | 30.4 | 33.5 | 31.9 | 31.5 | 35 | 30.8 | 35 | 30.8 | 31.3089 | 31.9164
| 1 | 5 | 31.6 | 33 | 38.8122 | 35.13 | 32.5658 | 31.3 | 30.01 | 31 | 31.9 | 30.3 | 33.7 | 32.2 | 29.8 | 35.2 | 33 | 35.2 | 29.7 | 30.9847 | 129
| 1 | 6 | 31.2 | 34 | 38.0209 | 39.55 | 29.8246 | 32 | 30.43 | 29.4 | 32.1 | 30.7 | 31.2 | 32 | 31.2 | 34.9 | 30.5 | 35.1 | 29.8 | 31.3439
| Mean |  | 31.5167 | 33 | 38.5452 | 37.9217 | 31.1302 | 32.2 | 30.165 | 30.4333 | 31.9333 | 30.6 | 32.6833 | 31.9 | 31.0667 | 35.2833 | 31.35 | 35.1333 | 30.45 | 31.8178
| Median |  | 31.4 | 33 | 38.675 | 38.22 | 30.9918 | 32.4 | 30.14 | 30.35 | 32 | 30.5 | 32.8 | 31.95 | 31.15 | 35.1 | 30.9 | 35.1 | 30.55 | 31.3264
| Std Dev. |  | 0.41191 | 0.63246 | 0.37878 | 1.5446 | 1.1933 | 0.4858 | 0.17237 | 0.83106 | 1.1183 | 0.37947 | 1.6179 | 0.43818 | 0.66533 | 0.4535 | 1.0154 | 0.10328 | 0.61563 | 1.0658
| Rel.Std.Dev. |  | 1.307% | 1.917% | 0.9827% | 4.073% | 3.833% | 1.509% | 0.5714% | 2.731% | 3.502% | 1.24% | 4.95% | 1.374% | 2.142% | 1.285% | 3.239% | 0.294% | 2.022% | 3.35%
| PDM3 |  | -1.253% | 3.395% | 20.77% | 18.82% | -2.463% | 0.8885% | -5.488% | -4.647% | 0.05294% | -4.125% | 2.403% | -0.0515% | -2.662% | 10.55% | -1.775% | 10.08% | -4.595% | -0.3091%
| Z-Score (Absolute) |  | 0.23 | 0.78 | 4.59 | 4.16 | 0.5 | 0.23 | 1.16 | 0.98 | 0.05 | 0.86 | 0.57 | 0.03 | 0.54 | 2.35 | 0.35 | 2.25 | 0.97 | 0.03
| Analytical results for Ti in OREAS 466 (Certified Value 0.587 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.6064 | 0.56083 | 0.58669 | 0.63028 | 0.6 | 0.58 | 0.6 | 0.59 | 0.58 | 0.5717 | 0.58718 | [7 empty] | 1
| 1 | 2 | 0.6062 | 0.59599 | 0.6037 | 0.61594 | 0.58 | 0.59 | 0.61 | 0.59 | 0.58 | 0.5733 | 0.59784 | [7 empty] | 0
| 1 | 3 | 0.6094 | 0.56865 | 0.57123 | 0.62246 | 0.59 | 0.58 | 0.56 | 0.59 | 0.57 | 0.5649 | 0.5877 | [7 empty] | 16
| 1 | 4 | 0.5874 | 0.55228 | 0.60359 | 0.62446 | 0.58 | 0.58 | 0.62 | 0.59 | 0.57 | 0.5671 | 0.59594 | [7 empty] | 0.58653
| 1 | 5 | 0.6153 | 0.55968 | 0.59263 | 0.61693 | 0.59 | 0.59 | 0.59 | 0.59 | 0.58 | 0.5634 | 0.5958 | [7 empty] | 130
| 1 | 6 | 0.6091 | 0.55409 | 0.59678 | 0.63151 | 0.6 | 0.59 | 0.61 | 0.59 | 0.58 | 0.5691 | 0.59771
| Mean |  | 0.60563 | 0.56526 | 0.59244 | 0.6236 | 0.59 | 0.585 | 0.59833 | 0.59 | 0.57667 | 0.56825 | 0.59369
| Median |  | 0.60775 | 0.56026 | 0.59471 | 0.62346 | 0.59 | 0.585 | 0.605 | 0.59 | 0.58 | 0.5681 | 0.59587
| Std Dev. |  | 0.0095185 | 0.016124 | 0.012272 | 0.0065158 | 0.0089443 | 0.0054772 | 0.02137 | 0 | 0.005164 | 0.0038521 | 0.0049215
| Rel.Std.Dev. |  | 1.572% | 2.852% | 2.072% | 1.045% | 1.516% | 0.9363% | 3.572% | 0% | 0.8955% | 0.6779% | 0.829%
| PDM3 |  | 3.258% | -3.627% | 1.008% | 6.32% | 0.5921% | -0.2603% | 2.013% | 0.5921% | -1.681% | -3.116% | 1.222%
| Z-Score (Absolute) |  | 1.27 | 2 | 0.2 | 2.72 | 0 | 0.4 | 0.67 | 0 | 1.08 | 1.76 | 0.3
| Analytical results for TiO2 in OREAS 466 (Certified Value 0.978 wt.%) | [19 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [7 empty] | wt.%
|  |  | PF*MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | [7 empty] | 3
3
| 1 | 1 | 1.0115 | 0.93553 | 0.97865 | 1.0514 | 1.001 | 0.967 | 1.001 | 0.984 | 0.967 | 0.95365 | 0.97947 | [7 empty] | 1
| 1 | 2 | 1.0112 | 0.99418 | 1.007 | 1.0275 | 0.967 | 0.984 | 1.018 | 0.984 | 0.967 | 0.95632 | 0.99725 | [7 empty] | 0
| 1 | 3 | 1.0165 | 0.94857 | 0.95287 | 1.0383 | 0.984 | 0.967 | 0.934 | 0.984 | 0.951 | 0.94231 | 0.98035 | [7 empty] | 16
| 1 | 4 | 0.97984 | 0.92126 | 1.0068 | 1.0417 | 0.967 | 0.967 | 1.034 | 0.984 | 0.951 | 0.94598 | 0.99409 | [7 empty] | 0.97839
| 1 | 5 | 1.0264 | 0.93361 | 0.98856 | 1.0291 | 0.984 | 0.984 | 0.984 | 0.984 | 0.967 | 0.93981 | 0.99385 | [7 empty] | 131
| 1 | 6 | 1.016 | 0.92427 | 0.99549 | 1.0534 | 1.001 | 0.984 | 1.018 | 0.984 | 0.967 | 0.94932 | 0.99704
| Mean |  | 1.0103 | 0.9429 | 0.98824 | 1.0402 | 0.984 | 0.9755 | 0.99817 | 0.984 | 0.96167 | 0.9479 | 0.99034
| Median |  | 1.0138 | 0.93457 | 0.99203 | 1.04 | 0.984 | 0.9755 | 1.0095 | 0.984 | 0.967 | 0.94765 | 0.99397
| Std Dev. |  | 0.015878 | 0.026896 | 0.020472 | 0.010869 | 0.015205 | 0.0093113 | 0.035768 | 0 | 0.0082624 | 0.0064257 | 0.0082096
| Rel.Std.Dev. |  | 1.572% | 2.852% | 2.072% | 1.045% | 1.545% | 0.9545% | 3.583% | 0% | 0.8592% | 0.6779% | 0.829%
| PDM3 |  | 3.258% | -3.627% | 1.008% | 6.32% | 0.5739% | -0.2949% | 2.022% | 0.5739% | -1.709% | -3.116% | 1.222%
| Z-Score (Absolute) |  | 1.27 | 2 | 0.2 | 2.72 | 0 | 0.4 | 0.67 | 0 | 1.08 | 1.76 | 0.3
| Analytical results for Tl in OREAS 466 (Indicative Value 0.77 ppm) | [19 empty] | Indicative
Tl | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.01: 2.07; 2.02: 2.17; 2.04: 2.23; 2.05: 2.24; 2.06: 2.25 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES/MS; 2.04: PF*MS; 2.05: PF*MS; 2.06: PF*MS | 2
2
| 1 | 1 | 2.01: 0.7; 2.02: 0.8; 2.04: 0.8; 2.05: < 1; 2.06: 0.8 | 1
| 1 | 2 | 2.01: 0.7; 2.02: 0.8; 2.04: 0.8; 2.05: < 1; 2.06: 0.8 | 6
| 1 | 3 | 2.01: 0.8; 2.02: 0.8; 2.04: 0.8; 2.05: < 1; 2.06: 0.8 | 16
| 1 | 4 | 2.01: 0.7; 2.02: 0.8; 2.04: 0.7; 2.05: < 1; 2.06: 0.8 | 0.77083
| 1 | 5 | 2.01: 0.8; 2.02: 0.8; 2.04: 0.7; 2.05: < 1; 2.06: 0.8 | 12
| 1 | 6 | 2.01: 0.7; 2.02: 0.8; 2.04: 0.8; 2.05: < 1; 2.06: 0.7
| Mean |  | 2.01: 0.73333; 2.02: 0.8; 2.04: 0.76667; 2.06: 0.78333
| Median |  | 2.01: 0.7; 2.02: 0.8; 2.04: 0.8; 2.06: 0.8
| Std Dev. |  | 2.01: 0.05164; 2.02: 1.2162e-16; 2.04: 0.05164; 2.06: 0.040825
| Rel.Std.Dev. |  | 2.01: 7.042%; 2.02: 1.52e-14%; 2.04: 6.736%; 2.06: 5.212%
| PDM3 |  | 2.01: -4.865%; 2.02: 3.784%; 2.04: -0.5405%; 2.06: 1.622%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.67; 2.04: 0; 2.05: 5.39; 2.06: 0.34
| Analytical results for Tm in OREAS 466 (Certified Value 0.51 ppm) | [19 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 0.6 | < 2 | 0.52425 | 0.5 | 0.62 | 0.49 | 0.5 | 0.4 | 0.5 | 0.52 | 0.49 | 0.51 | 0.5 | < 1 | 0.51 | < 1 |  |  | 1
| 1 | 2 | 0.6 | < 2 | < 1 | 0.5 | 0.65 | 0.52 | 0.41 | 0.4 | 0.52 | 0.49 | 0.45 | 0.52 | 0.5 | < 1 | 0.52 | < 1 |  |  | 0
| 1 | 3 | 0.5 | < 2 | 0.54033 | 0.5 | 0.67 | 0.51 | 0.49 | 0.42 | 0.47 | 0.51 | 0.49 | 0.51 | 0.5 | < 1 | 0.51 | < 1 |  |  | 16
| 1 | 4 | 0.5 | < 2 | 0.52216 | 0.6 | 0.65 | 0.49 | 0.61 | 0.37 | 0.46 | 0.56 | 0.5 | 0.53 | 0.5 | < 1 | 0.51 | < 1 |  |  | 0.50597
| 1 | 5 | 0.5 | < 2 | 0.51118 | 0.5 | 0.68 | 0.51 | 0.54 | 0.44 | 0.6 | 0.47 | 0.47 | 0.53 | 0.5 | < 1 | 0.47 | < 1 |  |  | 132
| 1 | 6 | 0.6 | < 2 | 0.54412 | 0.5 | 0.64 | 0.5 | 0.44 | 0.46 | 0.48 | 0.52 | 0.48 | 0.51 | 0.5 | < 1 | 0.51 | < 1
| Mean |  | 0.55 |  | 0.52841 | 0.51667 | 0.65167 | 0.50333 | 0.49833 | 0.415 | 0.505 | 0.51167 | 0.48 | 0.51833 | 0.5 |  | 0.505
| Median |  | 0.55 |  | 0.52425 | 0.5 | 0.65 | 0.505 | 0.495 | 0.41 | 0.49 | 0.515 | 0.485 | 0.515 | 0.5 |  | 0.51
| Std Dev. |  | 0.054772 |  | 0.013621 | 0.040825 | 0.02137 | 0.012111 | 0.071391 | 0.032094 | 0.051284 | 0.030605 | 0.017889 | 0.0098319 | 0 |  | 0.017607
| Rel.Std.Dev. |  | 9.959% |  | 2.578% | 7.902% | 3.279% | 2.406% | 14.33% | 7.733% | 10.16% | 5.981% | 3.727% | 1.897% | 0% |  | 3.486%
| PDM3 |  | 8.703% |  | 4.435% | 2.115% | 28.8% | -0.5207% | -1.509% | -17.98% | -0.1913% | 1.126% | -5.132% | 2.444% | -1.179% |  | -0.1913%
| Z-Score (Absolute) |  | NA | 50.15 | 1.97 | NA | 14.92 | 0.08 | 0.59 | 9.02 | 0.08 | 0.76 | 2.44 | 1.43 | 0.42 | 0.42 | 0.08 | 0.42
Results from laboratories 1 and 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm2O3 in OREAS 466 (Certified Value 0.58 ppm) | [19 empty] | Certified
Tm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 0.69 | < 2.3 | 0.59875 | 0.57 | 0.708 | 0.56 | 0.571 | 0.457 | 0.571 | 0.594 | 0.56 | 0.582 | 0.571 | < 1.1 | 0.582 | < 1.1 |  |  | 1
| 1 | 2 | 0.69 | < 2.3 | < 1.1 | 0.57 | 0.742 | 0.594 | 0.468 | 0.457 | 0.594 | 0.56 | 0.514 | 0.594 | 0.571 | < 1.1 | 0.594 | < 1.1 |  |  | 0
| 1 | 3 | 0.57 | < 2.3 | 0.61711 | 0.57 | 0.765 | 0.582 | 0.56 | 0.48 | 0.537 | 0.582 | 0.56 | 0.582 | 0.571 | < 1.1 | 0.582 | < 1.1 |  |  | 16
| 1 | 4 | 0.57 | < 2.3 | 0.59636 | 0.69 | 0.742 | 0.56 | 0.697 | 0.423 | 0.525 | 0.64 | 0.571 | 0.605 | 0.571 | < 1.1 | 0.582 | < 1.1 |  |  | 0.57612
| 1 | 5 | 0.57 | < 2.3 | 0.58382 | 0.57 | 0.777 | 0.582 | 0.617 | 0.503 | 0.685 | 0.537 | 0.537 | 0.605 | 0.571 | < 1.1 | 0.537 | < 1.1 |  |  | 133
| 1 | 6 | 0.69 | < 2.3 | 0.62144 | 0.57 | 0.731 | 0.571 | 0.503 | 0.525 | 0.548 | 0.594 | 0.548 | 0.582 | 0.571 | < 1.1 | 0.582 | < 1.1
| Mean |  | 0.63 |  | 0.6035 | 0.59 | 0.74417 | 0.57483 | 0.56933 | 0.47417 | 0.57667 | 0.5845 | 0.54833 | 0.59167 | 0.571 |  | 0.5765
| Median |  | 0.63 |  | 0.59875 | 0.57 | 0.742 | 0.5765 | 0.5655 | 0.4685 | 0.5595 | 0.588 | 0.554 | 0.588 | 0.571 |  | 0.582
| Std Dev. |  | 0.065727 |  | 0.015556 | 0.04899 | 0.024506 | 0.0136 | 0.081645 | 0.036488 | 0.058538 | 0.034996 | 0.020461 | 0.011325 | 1.2162e-16 |  | 0.019937
| Rel.Std.Dev. |  | 10.43% |  | 2.578% | 8.303% | 3.293% | 2.366% | 14.34% | 7.695% | 10.15% | 5.987% | 3.732% | 1.914% | 2.13e-14% |  | 3.458%
| PDM3 |  | 9.352% |  | 4.751% | 2.409% | 29.17% | -0.2237% | -1.178% | -17.7% | 0.09456% | 1.454% | -4.823% | 2.698% | -0.889% |  | 0.06563%
| Z-Score (Absolute) |  | 3.31 | 36.41 | 1.37 | 0.86 | 10.79 | 0.12 | 0.49 | 6.62 | 0 | 0.49 | 1.84 | 0.98 | 0.37 | 0.37 | 0 | 0.37
| Analytical results for U in OREAS 466 (Certified Value 3.44 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | 2
3
| 1 | 1 | 3.4 | 3 | 3.8634 | 4.933 | 3.9066 | 3.3 | 3.72 | 3.4 | 3.3 | 3.3 | 3.3 | 3.3 | 3.3 | 3.5 | 3.4 | 3.9 | 3.3 | 3.2864 | 1
| 1 | 2 | 3.3 | 3 | 3.8817 | 4.408 | 3.8493 | 3.5 | 3.71 | 3.5 | 3 | 3.4 | 3.7 | 3.5 | 3.2 | 3.8 | 3.8 | 3.9 | 3.23 | 3.611 | 0
| 1 | 3 | 3.4 | 3 | 3.8429 | 4.393 | 3.4094 | 3.2 | 3.71 | 3.4 | 3.1 | 3.5 | 3.2 | 3.2 | 3.3 | 3.8 | 3.4 | 3.7 | 3.24 | 3.245 | 16
| 1 | 4 | 3.5 | 3 | 3.8346 | 4.902 | 3.3197 | 3.4 | 3.67 | 3.3 | 3.2 | 3.4 | 3.3 | 3.7 | 3.4 | 3.7 | 3.4 | 3.7 | 3.24 | 3.3645 | 3.4449
| 1 | 5 | 3.3 | 3 | 3.8952 | 4.698 | 3.7901 | 3.6 | 3.76 | 3.4 | 3.1 | 3.4 | 3.3 | 3.6 | 3.3 | 3.7 | 3.6 | 3.8 | 3.18 | 3.2927 | 134
| 1 | 6 | 3.3 | 3 | 3.8724 | 4.46 | 3.808 | 3.5 | 3.77 | 3.3 | 3.1 | 3.4 | 3.3 | 3.2 | 3.3 | 3.5 | 3.1 | 3.7 | 3.19 | 3.4388
| Mean |  | 3.3667 | 3 | 3.865 | 4.6323 | 3.6805 | 3.4167 | 3.7233 | 3.3833 | 3.1333 | 3.4 | 3.35 | 3.4167 | 3.3 | 3.6667 | 3.45 | 3.7833 | 3.23 | 3.3731
| Median |  | 3.35 | 3 | 3.8679 | 4.579 | 3.7991 | 3.45 | 3.715 | 3.4 | 3.1 | 3.4 | 3.3 | 3.4 | 3.3 | 3.7 | 3.4 | 3.75 | 3.235 | 3.3286
| Std Dev. |  | 0.08165 | 0 | 0.023061 | 0.24688 | 0.24963 | 0.1472 | 0.036697 | 0.075277 | 0.10328 | 0.063246 | 0.17607 | 0.2137 | 0.063246 | 0.13663 | 0.23452 | 0.098319 | 0.042895 | 0.13519
| Rel.Std.Dev. |  | 2.425% | 0% | 0.5967% | 5.329% | 6.782% | 4.308% | 0.9856% | 2.225% | 3.296% | 1.86% | 5.256% | 6.255% | 1.917% | 3.726% | 6.798% | 2.599% | 1.328% | 4.008%
| PDM3 |  | -2.271% | -12.91% | 12.2% | 34.47% | 6.84% | -0.8197% | 8.082% | -1.787% | -9.044% | -1.304% | -2.755% | -0.8197% | -4.206% | 6.437% | 0.1479% | 9.824% | -6.238% | -2.086%
| Z-Score (Absolute) |  | 0.29 | NA | 2.59 | 7.03 | 1.53 | 0 | 1.77 | 0.19 | 1.64 | 0.1 | 0.39 | 0 | 0.67 | 1.44 | 0.19 | 2.12 | 1.08 | 0.25
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 466 (Certified Value 140 ppm) | [19 empty] | Certified
V | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.05; 2.05: 2.07; 2.06: 2.09; 2.07: 2.15; 2.08: 2.17; 2.09: 2.23; 2.10: 2.25; 2.11: 2.26 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*OES/MS; 2.06: BF*MS; 2.07: PF*OES; 2.08: PF*OES/MS; 2.09: PF*OES; 2.10: PF*MS; 2.11: BF*MS | 0
0
| 1 | 1 | 2.01: 133; 2.02: 132.0915; 2.04: 138.6327; 2.05: 140; 2.06: 143; 2.07: 130; 2.08: 140; 2.09: 140; 2.10: 145; 2.11: 140.2627 | 1
| 1 | 2 | 2.01: 133; 2.02: 137.402; 2.04: 148.7073; 2.05: 134; 2.06: 142; 2.07: 130; 2.08: 140; 2.09: 150; 2.10: 143; 2.11: 147.935 | 0
| 1 | 3 | 2.01: 134; 2.02: 134.3905; 2.04: 139.2305; 2.05: 130; 2.06: 144; 2.07: 130; 2.08: 140; 2.09: 140; 2.10: 144; 2.11: 141.8839 | 16
| 1 | 4 | 2.01: 132; 2.02: 130.5335; 2.04: 141.1868; 2.05: 131; 2.06: 152; 2.07: 140; 2.08: 140; 2.09: 145; 2.10: 143; 2.11: 140.1449 | 140.0062
| 1 | 5 | 2.01: 135; 2.02: 132.044; 2.04: 141.3585; 2.05: 135; 2.06: 136; 2.07: 130; 2.08: 140; 2.09: 145; 2.10: 142; 2.11: 140.5418 | 135
| 1 | 6 | 2.01: 134; 2.02: 130.125; 2.04: 141.9475; 2.05: 137; 2.06: 166; 2.07: 130; 2.08: 140; 2.09: 140; 2.10: 140; 2.11: 141.2638
| Mean |  | 2.01: 133.5; 2.02: 132.7644; 2.04: 141.8439; 2.05: 134.5; 2.06: 147.1667; 2.07: 131.6667; 2.08: 140; 2.09: 143.3333; 2.10: 142.8333; 2.11: 142.0053
| Median |  | 2.01: 133.5; 2.02: 132.0677; 2.04: 141.2726; 2.05: 134.5; 2.06: 143.5; 2.07: 130; 2.08: 140; 2.09: 142.5; 2.10: 143; 2.11: 140.9028
| Std Dev. |  | 2.01: 1.0488; 2.02: 2.7227; 2.04: 3.6035; 2.05: 3.7283; 2.06: 10.553; 2.07: 4.0825; 2.08: 0; 2.09: 4.0825; 2.10: 1.7224; 2.11: 2.9788
| Rel.Std.Dev. |  | 2.01: 0.7856%; 2.02: 2.051%; 2.04: 2.54%; 2.05: 2.772%; 2.06: 7.171%; 2.07: 3.101%; 2.08: 0%; 2.09: 2.848%; 2.10: 1.206%; 2.11: 2.098%
| PDM3 |  | 2.01: -4.647%; 2.02: -5.172%; 2.04: 1.313%; 2.05: -3.933%; 2.06: 5.114%; 2.07: -5.957%; 2.08: -0.004408%; 2.09: 2.376%; 2.10: 2.019%; 2.11: 1.428%
| Z-Score (Absolute) |  | 2.01: 1.71; 2.02: 1.86; 2.04: 0.02; 2.05: 1.51; 2.06: 1.06; 2.07: NA; 2.08: NA; 2.09: 0.29; 2.10: 0.18; 2.11: 0.02
Results from laboratories 15 and 17 were removed due to their 10 ppm reading resolution.
| Analytical results for V2O5 in OREAS 466 (Certified Value 250 ppm) | [19 empty] | Certified
V2O5 | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.05; 2.05: 2.07; 2.06: 2.09; 2.07: 2.15; 2.08: 2.17; 2.09: 2.23; 2.10: 2.25; 2.11: 2.26 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*OES/MS; 2.06: BF*MS; 2.07: PF*OES; 2.08: PF*OES/MS; 2.09: PF*OES; 2.10: PF*MS; 2.11: BF*MS | 0
0
| 1 | 1 | 2.01: 237.4; 2.02: 235.8098; 2.04: 247.4871; 2.05: 249.9; 2.06: 255.3; 2.07: 232; 2.08: 250; 2.09: 249.9; 2.10: 258.85; 2.11: 250.3969 | 1
| 1 | 2 | 2.01: 237.4; 2.02: 245.2901; 2.04: 265.4723; 2.05: 239.2; 2.06: 253.5; 2.07: 232; 2.08: 250; 2.09: 267.8; 2.10: 255.28; 2.11: 264.0935 | 13
| 1 | 3 | 2.01: 239.2; 2.02: 239.9139; 2.04: 248.5543; 2.05: 232.1; 2.06: 257.1; 2.07: 232; 2.08: 250; 2.09: 249.9; 2.10: 257.07; 2.11: 253.2911 | 16
| 1 | 4 | 2.01: 235.6; 2.02: 233.0284; 2.04: 252.0467; 2.05: 233.9; 2.06: 271.4; 2.07: 250; 2.08: 250; 2.09: 258.9; 2.10: 255.28; 2.11: 250.1867 | 249.939
| 1 | 5 | 2.01: 241; 2.02: 235.725; 2.04: 252.3532; 2.05: 241; 2.06: 242.8; 2.07: 232; 2.08: 250; 2.09: 258.9; 2.10: 253.5; 2.11: 250.8952 | 136
| 1 | 6 | 2.01: 239.2; 2.02: 232.2991; 2.04: 253.4047; 2.05: 244.6; 2.06: 296.3; 2.07: 232; 2.08: 250; 2.09: 249.9; 2.10: 249.93; 2.11: 252.1841
| Mean |  | 2.01: 238.3; 2.02: 237.011; 2.04: 253.2197; 2.05: 240.1167; 2.06: 262.7333; 2.07: 235; 2.08: 250; 2.09: 255.8833; 2.10: 254.985; 2.11: 253.5079
| Median |  | 2.01: 238.3; 2.02: 235.7673; 2.04: 252.1999; 2.05: 240.1; 2.06: 256.2; 2.07: 232; 2.08: 250; 2.09: 254.4; 2.10: 255.28; 2.11: 251.5396
| Std Dev. |  | 2.01: 1.8879; 2.02: 4.8606; 2.04: 6.433; 2.05: 6.6385; 2.06: 18.8224; 2.07: 7.3485; 2.08: 0; 2.09: 7.3159; 2.10: 3.073; 2.11: 5.3177
| Rel.Std.Dev. |  | 2.01: 0.7922%; 2.02: 2.051%; 2.04: 2.54%; 2.05: 2.765%; 2.06: 7.164%; 2.07: 3.127%; 2.08: 0%; 2.09: 2.859%; 2.10: 1.205%; 2.11: 2.098%
| PDM3 |  | 2.01: -4.657%; 2.02: -5.172%; 2.04: 1.313%; 2.05: -3.93%; 2.06: 5.119%; 2.07: -5.977%; 2.08: 0.0244%; 2.09: 2.378%; 2.10: 2.019%; 2.11: 1.428%
| Z-Score (Absolute) |  | 2.01: 1.71; 2.02: 1.86; 2.04: 0.02; 2.05: 1.51; 2.06: 1.06; 2.07: NA; 2.08: NA; 2.09: 0.29; 2.10: 0.18; 2.11: 0.02
Results from laboratories 15 and 17 were removed due to their 10 ppm reading resolution.
| Analytical results for W in OREAS 466 (Certified Value 94 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 |  |  |  | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | BF*MS | BF*MS |  |  |  | 0
1
| 1 | 1 | 86 | 13.2488 | 93.838 | 75 | 94.3 | 97 | 89 | 95 | 96 | 96 | < 100 | 1090000 | 95 | 94 | 97.2392 |  |  |  | 1
| 1 | 2 | 83 | 16.274 | 99.776 | 70.7 | 97.1 | 88 | 86 | 99 | 101 | 97 | < 100 | 1140000 | 91 | 94.9 | 100.7156 |  |  |  | 0
| 1 | 3 | 85 | 13.1467 | 99.481 | 76.8 | 95.8 | 93 | 88 | 89 | 99 | 99 | < 100 | 1100000 | 97 | 95.3 | 96.0445 |  |  |  | 16
| 1 | 4 | 80 | 13.4185 | 91.5377 | 77.1 | 93.7 | 98 | 88 | 100 | 103 | 100 | < 100 | 1140000 | 98 | 95.2 | 100.1846 |  |  |  | 94.4442
| 1 | 5 | 85 | 13.1893 | 93.7705 | 73.7 | 93.9 | 99 | 89 | 94 | 98 | 100 | < 100 | 1130000 | 95 | 94.6 | 97.7398 |  |  |  | 137
| 1 | 6 | 89 | 13.4743 | 93.284 | 75.7 | 92.2 | 101 | 89 | 87 | 99 | 102 | < 100 | 1120000 | 96 | 92.7 | 97.006
| Mean |  | 84.6667 | 13.7919 | 95.2812 | 74.8333 | 94.5 | 96 | 88.1667 | 94 | 99.3333 | 99 |  | 1120000 | 95.3333 | 94.45 | 98.155
| Median |  | 85 | 13.3337 | 93.8042 | 75.35 | 94.1 | 97.5 | 88.5 | 94.5 | 99 | 99.5 |  | 1125000 | 95.5 | 94.75 | 97.4895
| Std Dev. |  | 3.0111 | 1.2228 | 3.4701 | 2.3729 | 1.7193 | 4.7329 | 1.169 | 5.2154 | 2.4221 | 2.1909 |  | 20976.177 | 2.4221 | 0.97724 | 1.8688
| Rel.Std.Dev. |  | 3.556% | 8.866% | 3.642% | 3.171% | 1.819% | 4.93% | 1.326% | 5.548% | 2.438% | 2.213% |  | 1.873% | 2.541% | 1.035% | 1.904%
| PDM3 |  | -10.35% | -85.4% | 0.8862% | -20.76% | 0.05909% | 1.647% | -6.647% | -0.4703% | 5.177% | 4.824% |  | 1185785% | 0.9414% | 0.006146% | 3.929%
| Z-Score (Absolute) |  | 1.61 | 13.26 | 0.13 | 3.23 | 0 | 0.25 | 1.04 | 0.08 | 0.8 | 0.74 | 7.31 | NA | 0.14 | 0 | 0.6
| Analytical results for Y in OREAS 466 (Certified Value 34.6 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  | 1
2
| 1 | 1 | 35 | 32 | 34.5592 | 40.4926 | 34.1 | 40.79 | 33.7 | 37 | 31 | 35 | 35 | 37 | 33.6 | 35.7 | 24.7 | 35.78 | 32.0344 |  | 1
| 1 | 2 | 34.4 | 33 | 36.4044 | 35.6142 | 32.5 | 40.83 | 33.7 | 34 | 31 | 38 | 35 | 36 | 34.4 | 35 | 24.3 | 34.63 | 34.7017 |  | 0
| 1 | 3 | 34.1 | 32 | 35.1366 | 40.3551 | 34 | 40.12 | 33.4 | 37 | 33 | 35 | 34 | 35 | 34.2 | 34.6 | 24.6 | 34.67 | 32.3212 |  | 16
| 1 | 4 | 34.7 | 32 | 35.2724 | 40.3669 | 33.5 | 40.01 | 33.6 | 38 | 32 | 36 | 34 | 36 | 34.7 | 36 | 25 | 34.67 | 34.2602 |  | 34.6308
| 1 | 5 | 34.1 | 31 | 34.8041 | 37.7872 | 32.3 | 40.59 | 33.1 | 40 | 33 | 36 | 35 | 36 | 34.9 | 34 | 24.5 | 33.81 | 32.2184 |  | 138
| 1 | 6 | 35 | 33 | 35.1001 | 38.368 | 33.4 | 40.97 | 32 | 35 | 33 | 35 | 35 | 35 | 34.4 | 34.5 | 24.5 | 33.71 | 32.1078
| Mean |  | 34.55 | 32.1667 | 35.2128 | 38.8307 | 33.3 | 40.5517 | 33.25 | 36.8333 | 32.1667 | 35.8333 | 34.6667 | 35.8333 | 34.3667 | 34.9667 | 24.6 | 34.545 | 32.9406
| Median |  | 34.55 | 32 | 35.1184 | 39.3615 | 33.45 | 40.69 | 33.5 | 37 | 32.5 | 35.5 | 35 | 36 | 34.4 | 34.8 | 24.55 | 34.65 | 32.2698
| Std Dev. |  | 0.41352 | 0.75277 | 0.63815 | 1.9542 | 0.751 | 0.39761 | 0.65345 | 2.137 | 0.98319 | 1.169 | 0.5164 | 0.75277 | 0.45019 | 0.7607 | 0.23664 | 0.74849 | 1.2052
| Rel.Std.Dev. |  | 1.197% | 2.34% | 1.812% | 5.033% | 2.255% | 0.9805% | 1.965% | 5.802% | 3.057% | 3.262% | 1.49% | 2.101% | 1.31% | 2.176% | 0.962% | 2.167% | 3.659%
| PDM3 |  | -0.2334% | -7.116% | 1.681% | 12.13% | -3.843% | 17.1% | -3.987% | 6.36% | -7.116% | 3.472% | 0.1035% | 3.472% | -0.7628% | 0.9698% | -28.97% | -0.2478% | -4.881%
| Z-Score (Absolute) |  | 0 | 1.25 | 0.35 | 2.25 | 0.66 | 3.15 | 0.68 | 1.2 | 1.25 | 0.67 | 0.06 | 0.67 | 0.1 | 0.22 | 5.23 | 0 | 0.85
| Analytical results for Y2O3 in OREAS 466 (Certified Value 44 ppm) | [19 empty] | Certified
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS |  | 1
2
| 1 | 1 | 44.45 | 40.6 | 43.8868 | 51.4216 | 43.3 | 51.799 | 42.8 | 47 | 39.4 | 44.4 | 44.4 | 47 | 42.67 | 45.335 | 31.37 | 45.437 | 40.6805 |  | 1
| 1 | 2 | 43.68 | 41.9 | 46.23 | 45.2265 | 41.27 | 51.85 | 42.8 | 43.2 | 39.4 | 48.3 | 44.4 | 45.7 | 43.68 | 44.446 | 30.86 | 43.9766 | 44.0676 |  | 15
| 1 | 3 | 43.3 | 40.6 | 44.62 | 51.2469 | 43.18 | 50.948 | 42.41 | 47 | 41.9 | 44.4 | 43.2 | 44.4 | 43.43 | 43.939 | 31.24 | 44.0274 | 41.0447 |  | 16
| 1 | 4 | 44.07 | 40.6 | 44.7924 | 51.262 | 42.54 | 50.809 | 42.67 | 48.3 | 40.6 | 45.7 | 43.2 | 45.7 | 44.07 | 45.716 | 31.75 | 44.0274 | 43.5071 |  | 43.9777
| 1 | 5 | 43.3 | 39.4 | 44.1977 | 47.986 | 41.02 | 51.545 | 42.03 | 50.8 | 41.9 | 45.7 | 44.4 | 45.7 | 44.32 | 43.177 | 31.11 | 42.9353 | 40.9142 |  | 139
| 1 | 6 | 44.45 | 41.9 | 44.5737 | 48.7235 | 42.41 | 52.028 | 40.64 | 44.4 | 41.9 | 44.4 | 44.4 | 44.4 | 43.68 | 43.812 | 31.11 | 42.8083 | 40.7737
| Mean |  | 43.875 | 40.8333 | 44.7168 | 49.3111 | 42.2867 | 51.4965 | 42.225 | 46.7833 | 40.85 | 45.4833 | 44 | 45.4833 | 43.6417 | 44.4042 | 31.24 | 43.8687 | 41.8313
| Median |  | 43.875 | 40.6 | 44.5968 | 49.9852 | 42.475 | 51.672 | 42.54 | 47 | 41.25 | 45.05 | 44.4 | 45.7 | 43.68 | 44.1925 | 31.175 | 44.002 | 40.9795
| Std Dev. |  | 0.52888 | 0.94798 | 0.81039 | 2.4816 | 0.95314 | 0.50494 | 0.82991 | 2.7221 | 1.2309 | 1.5198 | 0.61968 | 0.9786 | 0.57192 | 0.96562 | 0.30173 | 0.9505 | 1.5305
| Rel.Std.Dev. |  | 1.205% | 2.322% | 1.812% | 5.033% | 2.254% | 0.9805% | 1.965% | 5.818% | 3.013% | 3.341% | 1.408% | 2.152% | 1.31% | 2.175% | 0.9658% | 2.167% | 3.659%
| PDM3 |  | -0.2335% | -7.15% | 1.681% | 12.13% | -3.845% | 17.1% | -3.985% | 6.38% | -7.112% | 3.424% | 0.05072% | 3.424% | -0.7641% | 0.9697% | -28.96% | -0.2479% | -4.881%
| Z-Score (Absolute) |  | 0 | 1.25 | 0.35 | 2.25 | 0.66 | 3.15 | 0.68 | 1.2 | 1.25 | 0.67 | 0.06 | 0.67 | 0.1 | 0.22 | 5.23 | 0 | 0.85
| Analytical results for Yb in OREAS 466 (Certified Value 3.29 ppm) | [19 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 3.5 | 3 | 3.0793 | 3.6 | 2.89 | 3.28 | 3.7 | 3.3 | 3.3 | 3.4 | 3.3 | 3.4 | 3.4 | 2.7 | 3.2 | 3.4053 |  |  | 1
| 1 | 2 | 3.3 | 3 | 2.9396 | 3.3 | 3.35 | 3.32 | 3.2 | 3.2 | 3.9 | 3.3 | 3.2 | 3.4 | 3.2 | 2.5 | 3.5 | 3.4658 |  |  | 0
| 1 | 3 | 3.3 | 3 | 3.2015 | 3.8 | 2.69 | 3.18 | 3.5 | 3.3 | 3.4 | 2.7 | 3.1 | 3.4 | 3.4 | 2.7 | 3.4 | 3.3086 |  |  | 16
| 1 | 4 | 3.6 | 3 | 3.2014 | 3.3 | 3.27 | 3.14 | 3.2 | 3.2 | 3.7 | 3.4 | 3.3 | 3.3 | 3.2 | 2.7 | 3.3 | 3.2667 |  |  | 3.2944
| 1 | 5 | 3.2 | 3 | 3.0558 | 3.1 | 2.95 | 3.25 | 3.1 | 3.5 | 3.4 | 2.9 | 3.2 | 3.3 | 3 | 2.6 | 3.1 | 3.0278 |  |  | 140
| 1 | 6 | 3.5 | 3 | 3.1665 | 3.7 | 2.74 | 3.2 | 3.2 | 3.5 | 3.5 | 3.1 | 3.2 | 3.3 | 3.5 | 2.6 | 3.2 | 3.4849
| Mean |  | 3.4 | 3 | 3.1074 | 3.4667 | 2.9817 | 3.2283 | 3.3167 | 3.3333 | 3.5333 | 3.1333 | 3.2167 | 3.35 | 3.2833 | 2.6333 | 3.2833 | 3.3265
| Median |  | 3.4 | 3 | 3.1229 | 3.45 | 2.92 | 3.225 | 3.2 | 3.3 | 3.45 | 3.2 | 3.2 | 3.35 | 3.3 | 2.65 | 3.25 | 3.3569
| Std Dev. |  | 0.15492 | 0 | 0.10277 | 0.27325 | 0.27265 | 0.067057 | 0.23166 | 0.13663 | 0.22509 | 0.28752 | 0.075277 | 0.054772 | 0.18348 | 0.08165 | 0.1472 | 0.16958
| Rel.Std.Dev. |  | 4.556% | 0% | 3.307% | 7.882% | 9.144% | 2.077% | 6.985% | 4.099% | 6.371% | 9.176% | 2.34% | 1.635% | 5.588% | 3.101% | 4.483% | 5.098%
| PDM3 |  | 3.207% | -8.935% | -5.676% | 5.231% | -9.492% | -2.004% | 0.6774% | 1.183% | 7.254% | -4.888% | -2.358% | 1.689% | -0.3344% | -20.07% | -0.3344% | 0.9761%
| Z-Score (Absolute) |  | 1.18 | NA | 1.78 | 1.85 | 3.05 | 0.56 | 0.34 | 0.51 | 2.53 | 1.52 | 0.67 | 0.67 | 0 | 6.57 | 0 | 0.44
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb2O3 in OREAS 466 (Certified Value 3.72 ppm) | [19 empty] | Certified
Yb2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.23 | 2.24 | 2.25 | 2.26 |  |  | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS |  |  | 2
3
| 1 | 1 | 3.99 | 3.4 | 3.5064 | 4.1 | 3.291 | 3.735 | 4.21 | 3.76 | 3.76 | 3.87 | 3.76 | 3.87 | 3.87 | 3.07 | 3.64 | 3.8776 |  |  | 1
| 1 | 2 | 3.76 | 3.4 | 3.3473 | 3.76 | 3.815 | 3.78 | 3.64 | 3.64 | 4.44 | 3.76 | 3.64 | 3.87 | 3.64 | 2.85 | 3.99 | 3.9465 |  |  | 0
| 1 | 3 | 3.76 | 3.4 | 3.6456 | 4.33 | 3.063 | 3.621 | 3.99 | 3.76 | 3.87 | 3.07 | 3.53 | 3.87 | 3.87 | 3.07 | 3.87 | 3.7675 |  |  | 16
| 1 | 4 | 4.1 | 3.4 | 3.6455 | 3.76 | 3.724 | 3.576 | 3.64 | 3.64 | 4.21 | 3.87 | 3.76 | 3.76 | 3.64 | 3.07 | 3.76 | 3.7198 |  |  | 3.7168
| 1 | 5 | 3.64 | 3.4 | 3.4797 | 3.53 | 3.359 | 3.701 | 3.53 | 3.99 | 3.87 | 3.3 | 3.64 | 3.76 | 3.42 | 2.96 | 3.53 | 3.4478 |  |  | 141
| 1 | 6 | 3.99 | 3.4 | 3.6057 | 4.21 | 3.12 | 3.644 | 3.64 | 3.99 | 3.99 | 3.53 | 3.64 | 3.76 | 3.99 | 2.96 | 3.64 | 3.9682
| Mean |  | 3.8733 | 3.4 | 3.5384 | 3.9483 | 3.3953 | 3.6762 | 3.775 | 3.7967 | 4.0233 | 3.5667 | 3.6617 | 3.815 | 3.7383 | 2.9967 | 3.7383 | 3.7879
| Median |  | 3.875 | 3.4 | 3.5561 | 3.93 | 3.325 | 3.6725 | 3.64 | 3.76 | 3.93 | 3.645 | 3.64 | 3.815 | 3.755 | 3.015 | 3.7 | 3.8225
| Std Dev. |  | 0.17818 | 0 | 0.11702 | 0.31083 | 0.31068 | 0.076119 | 0.26463 | 0.15908 | 0.25532 | 0.3288 | 0.087274 | 0.060249 | 0.20913 | 0.089815 | 0.16964 | 0.1931
| Rel.Std.Dev. |  | 4.6% | 0% | 3.307% | 7.872% | 9.15% | 2.071% | 7.01% | 4.19% | 6.346% | 9.219% | 2.383% | 1.579% | 5.594% | 2.997% | 4.538% | 5.098%
| PDM3 |  | 4.213% | -8.522% | -4.8% | 6.231% | -8.648% | -1.092% | 1.567% | 2.15% | 8.248% | -4.038% | -1.482% | 2.643% | 0.5805% | -19.37% | 0.5805% | 1.914%
| Z-Score (Absolute) |  | 0.86 | 2.08 | 1.29 | 1.35 | 2.22 | 0.4 | 0.25 | 0.37 | 1.84 | 1.1 | 0.49 | 0.49 | 0 | 4.78 | 0 | 0.32
| Analytical results for Zn in OREAS 466 (Certified Value 47.4 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.15 | 2.17 | 2.23 | 2.25 | 2.26 | [8 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES | [8 empty] | 1
2
| 1 | 1 | 46 | 45.1836 | 77.206 | 50 | < 100 | 50 | 50 | 50 | 50 | < 100 | [8 empty] | 1
| 1 | 2 | 46 | 43.6466 | 100.5183 | 40 | < 100 | < 50 | 50 | 50 | 50 | < 100 | [8 empty] | 0
| 1 | 3 | 46 | 47.254 | 134.9715 | 50 | < 100 | < 50 | 40 | 50 | 50 | < 100 | [8 empty] | 16
| 1 | 4 | 49 | 43.6496 | 80.9393 | 40 | < 100 | < 50 | 50 | 50 | 50 | < 100 | [8 empty] | 47.4348
| 1 | 5 | 46 | 45.3698 | 78.953 | 40 | < 100 | < 50 | 40 | 45 | 52 | < 100 | [8 empty] | 142
| 1 | 6 | 47 | 44.157 | 74.7402 | 50 | < 100 | < 50 | 40 | 55 | 51 | < 100
| Mean |  | 2.01: 46.6667; 2.02: 44.8768; 2.06: 91.2214; 2.07: 45; 2.09: 50; 2.10: 45; 2.11: 50; 2.12: 50.5
| Median |  | 2.01: 46; 2.02: 44.6703; 2.06: 79.9461; 2.07: 45; 2.09: 50; 2.10: 45; 2.11: 50; 2.12: 50
| Std Dev. |  | 2.01: 1.2111; 2.02: 1.3803; 2.06: 23.3444; 2.07: 5.4772; 2.10: 5.4772; 2.11: 3.1623; 2.12: 0.83666
| Rel.Std.Dev. |  | 2.01: 2.595%; 2.02: 3.076%; 2.06: 25.59%; 2.07: 12.17%; 2.10: 12.17%; 2.11: 6.325%; 2.12: 1.657%
| PDM3 |  | 2.01: -1.619%; 2.02: -5.393%; 2.06: 92.31%; 2.07: -5.133%; 2.09: 5.408%; 2.10: -5.133%; 2.11: 5.408%; 2.12: 6.462%
| Z-Score (Absolute) |  | 0.41 | 0.85 | 10.52 | 0.82 | 0.41 | 4.7 | 0.82 | 0.41 | 0.53 | 0.41
| Analytical results for Zr in OREAS 466 (Certified Value 283 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.06 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.25 | 2.26 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | [7 empty] | 0
0
| 1 | 1 | 284 | 189 | 286.2275 | 276.8 | 280 | 270 | 280 | 300 | 300 | 268.2 | 262.7802 | [7 empty] | 1
| 1 | 2 | 277 | 188 | 289.0212 | 288.9 | 250 | 270 | 310 | 290 | 290 | 281.6 | 280.2336 | [7 empty] | 0
| 1 | 3 | 278 | 185 | 299.6656 | 285.9 | 270 | 280 | 280 | 280 | 300 | 269.5 | 267.9161 | [7 empty] | 16
| 1 | 4 | 279 | 184 | 291.4503 | 288.3 | 280 | 270 | 300 | 290 | 300 | 312.7 | 273.1777 | [7 empty] | 282.6915
| 1 | 5 | 273 | 183 | 286.6789 | 284.4 | 270 | 270 | 290 | 290 | 300 | 360.8 | 273.1537 | [7 empty] | 143
| 1 | 6 | 279 | 190 | 289.4328 | 284.9 | 260 | 290 | 280 | 290 | 290 | 340 | 264.5546
| Mean |  | 278.3333 | 186.5 | 290.4127 | 284.8667 | 268.3333 | 275 | 290 | 290 | 296.6667 | 305.4667 | 270.3027
| Median |  | 278.5 | 186.5 | 289.227 | 285.4 | 270 | 270 | 285 | 290 | 300 | 297.15 | 270.5349
| Std Dev. |  | 3.559 | 2.881 | 4.9203 | 4.3454 | 11.6905 | 8.3666 | 12.6491 | 6.3246 | 5.164 | 38.881 | 6.486
| Rel.Std.Dev. |  | 1.279% | 1.545% | 1.694% | 1.525% | 4.357% | 3.042% | 4.362% | 2.181% | 1.741% | 12.73% | 2.4%
| PDM3 |  | -1.542% | -34.03% | 2.731% | 0.7694% | -5.079% | -2.721% | 2.585% | 2.585% | 4.944% | 8.057% | -4.382%
| Z-Score (Absolute) |  | 0.45 | 6.72 | 0.38 | 0 | 1.13 | 0.67 | 0.35 | 0.35 | 0.81 | 1.41 | 1
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 466 (Certified Value 15.81 wt.%) | [13 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 16.09 | 15.59 | 16.112 | 30.5532 | 15.94 | 15.46 | 16.09 | 15.04 | 15.7248 | 15.63 | 15.89 | 15.9 | 1
| 1 | 2 | 16.02 | 15.64 | 16.116 | 31.2221 | 15.86 | 15.45 | 16.1 | 15.45 | 15.7181 | 15.63 | 15.97 | 16.1 | 0
| 1 | 3 | 15.96 | 15.84 | 16.183 | 30.7762 | 15.93 | 15.36 | 16.11 | 15.305 | 15.7317 | 15.55 | 15.83 | 16 | 16
| 1 | 4 | 16.06 | 16.07 | 16.198 | 30.9802 | 15.87 | 15.43 | 16.1 | 15.4 | 15.7324 | 15.5 | 15.87 | 15.8 | 15.8094
| 1 | 5 | 16.04 | 15.67 | 16.111 | 30.78 | 15.7 | 15.44 | 16.07 | 15.4 | 15.8667 | 15.59 | 16.03 | 16 | 145
| 1 | 6 | 16.05 | 15.64 | 16.108 | 30.7082 | 15.79 | 15.42 | 16.1 | 15.465 | 15.738 | 15.57 | 15.91 | 16
| Mean |  | 16.0367 | 15.7417 | 16.138 | 30.8366 | 15.8483 | 15.4267 | 16.095 | 15.3433 | 15.7519 | 15.5783 | 15.9167 | 15.9667
| Median |  | 16.045 | 15.655 | 16.114 | 30.7781 | 15.865 | 15.435 | 16.1 | 15.4 | 15.732 | 15.58 | 15.9 | 16
| Std Dev. |  | 0.044121 | 0.18236 | 0.041022 | 0.23362 | 0.090646 | 0.03559 | 0.013784 | 0.1588 | 0.056649 | 0.049967 | 0.072296 | 0.10328
| Rel.Std.Dev. |  | 0.2751% | 1.158% | 0.2542% | 0.7576% | 0.572% | 0.2307% | 0.08564% | 1.035% | 0.3596% | 0.3207% | 0.4542% | 0.6468%
| PDM3 |  | 1.437% | -0.4288% | 2.078% | 95.05% | 0.246% | -2.421% | 1.806% | -2.948% | -0.3637% | -1.462% | 0.6782% | 0.9944%
| Z-Score (Absolute) |  | 0.57 | 0.52 | 0.94 | 55 | 0.13 | 1.68 | 0.78 | 1.98 | 0.48 | 1.12 | 0.13 | 0.31
| Analytical results for As2O3 in OREAS 466 (Indicative Value 600 ppm) | [13 empty] | Indicative
As2O3 | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.23 | [11 empty] | ppm
|  |  | BF*XRF | [11 empty] | 0
0
| 1 | 1 | 600 | [11 empty] | 1
| 1 | 2 | 600 | [11 empty] | 9
| 1 | 3 | 600 | [11 empty] | 16
| 1 | 4 | 600 | [11 empty] | 600
| 1 | 5 | 600 | [11 empty] | 15
| 1 | 6 | 600
| Mean |  | 600
| Median |  | 600
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 466 (Certified Value 755 ppm) | [13 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.18 | 2.23 | 2.25 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 900 | 710 | 1284 | 700 | 780 | 787.8585 | 800 | 700 |  |  |  |  | 1
| 1 | 2 | 800 | 710 | 1273 | 700 | 780 | 806.4593 | 800 | 600 |  |  |  |  | 0
| 1 | 3 | 900 | 650 | 1239 | 800 | 780 | 780.2313 | 700 | 600 |  |  |  |  | 16
| 1 | 4 | 700 | 670 | 1262 | 700 | 780 | 764.4337 | 700 | 500 |  |  |  |  | 754.9696
| 1 | 5 | 900 | 660 | 1284 | 700 | 780 | 786.4904 | 800 | 600 |  |  |  |  | 146
| 1 | 6 | 600 | 670 | 1250 | 700 | 780 | 794.134 | 800 | 700
| Mean |  | 800 | 678.3333 | 1265.3333 | 716.6667 | 780 | 786.6012 | 766.6667 | 616.6667
| Median |  | 850 | 670 | 1267.5 | 700 | 780 | 787.1745 | 800 | 600
| Std Dev. |  | 126.4911 | 25.6255 | 18.4138 | 40.8248 | 0 | 14.0332 | 51.6398 | 75.2773
| Rel.Std.Dev. |  | 15.81% | 3.778% | 1.455% | 5.696% | 0% | 1.784% | 6.736% | 12.21%
| PDM3 |  | 5.965% | -10.15% | 67.6% | -5.073% | 3.315% | 4.19% | 1.549% | -18.32%
| Z-Score (Absolute) |  | 0.42 | 1.55 | 7.95 | 0.93 | 0.12 | 0.2 | 0.12 | 2.55
| Analytical results for Bi2O3 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Bi2O3 | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.07 | 2.23 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 170 | < 100 | [10 empty] | 1
| 1 | 2 | < 170 | < 100 | [10 empty] | 10
| 1 | 3 | < 170 | < 100 | [10 empty] | 16
| 1 | 4 | < 170 | < 100 | [10 empty] | < 100
| 1 | 5 | < 170 | < 100 | [10 empty] | 16
| 1 | 6 | < 170 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for CaO in OREAS 466 (Certified Value 0.318 wt.%) | [13 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.33 | 0.21 | 0.3 | 0.6296 | 0.31 | 0.29 | 0.34 | 0.335 | 0.32652 | 0.3 | 0.32 | 0.36 | 1
| 1 | 2 | 0.33 | 0.21 | 0.3 | 0.4645 | 0.3 | 0.27 | 0.34 | 0.335 | 0.32593 | 0.3 | 0.32 | 0.34 | 0
| 1 | 3 | 0.33 | 0.24 | 0.3 | 0.4771 | 0.31 | 0.3 | 0.34 | 0.325 | 0.32971 | 0.31 | 0.32 | 0.33 | 16
| 1 | 4 | 0.33 | 0.23 | 0.3 | 0.6268 | 0.32 | 0.27 | 0.34 | 0.315 | 0.32864 | 0.3 | 0.32 | 0.33 | 0.31843
| 1 | 5 | 0.33 | 0.22 | 0.303 | 0.4687 | 0.32 | 0.27 | 0.34 | 0.335 | 0.32952 | 0.31 | 0.33 | 0.34 | 147
| 1 | 6 | 0.33 | 0.21 | 0.303 | 0.4603 | 0.31 | 0.28 | 0.34 | 0.32 | 0.32426 | 0.3 | 0.32 | 0.35
| Mean |  | 0.33 | 0.22 | 0.301 | 0.52117 | 0.31167 | 0.28 | 0.34 | 0.3275 | 0.32743 | 0.30333 | 0.32167 | 0.34167
| Median |  | 0.33 | 0.215 | 0.3 | 0.4729 | 0.31 | 0.275 | 0.34 | 0.33 | 0.32758 | 0.3 | 0.32 | 0.34
| Std Dev. |  | 0 | 0.012649 | 0.0015492 | 0.083098 | 0.0075277 | 0.012649 | 0 | 0.0088034 | 0.0021989 | 0.005164 | 0.0040825 | 0.01169
| Rel.Std.Dev. |  | 0% | 5.75% | 0.5147% | 15.94% | 2.415% | 4.518% | 0% | 2.688% | 0.6716% | 1.702% | 1.269% | 3.422%
| PDM3 |  | 3.635% | -30.91% | -5.473% | 63.67% | -2.123% | -12.07% | 6.775% | 2.85% | 2.827% | -4.74% | 1.018% | 7.299%
| Z-Score (Absolute) |  | 0.23 | 4.33 | 0.98 | 8.14 | 0.53 | 1.84 | 0.64 | 0.12 | 0.12 | 0.88 | 0.12 | 0.71
| Analytical results for CeO2 in OREAS 466 (Certified Value 0.044 wt.%) | [13 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.05 | 0.043 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.049 | 0.031 | 0.051239 | 0.04 | 1
| 1 | 2 | 0.05 | 0.0467 | 0.05 | 0.05 | 0.03 | 0.03 | 0.04 | 0.04 | 0.049 | 0.052 | 0.051353 | 0.04 | 0
| 1 | 3 | 0.04 | 0.0418 | 0.06 | 0.05 | 0.04 | 0.03 | 0.04 | 0.04 | 0.049 | 0.031 | 0.050234 | 0.04 | 16
| 1 | 4 | 0.05 | 0.0369 | 0.06 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.049 | 0.042 | 0.049948 | 0.04 | 0.04363
| 1 | 5 | 0.04 | 0.0418 | 0.05 | 0.05 | 0.03 | 0.04 | 0.04 | 0.04 | 0.061 | 0.042 | 0.049826 | 0.04 | 148
| 1 | 6 | 0.04 | 0.0405 | 0.06 | 0.05 | 0.03 | 0.03 | 0.04 | 0.04 | 0.049 | 0.042 | 0.049824 | 0.04
| Mean |  | 0.045 | 0.041783 | 0.055 | 0.05 | 0.035 | 0.035 | 0.04 | 0.04 | 0.051 | 0.04 | 0.050404 | 0.04
| Median |  | 0.045 | 0.0418 | 0.055 | 0.05 | 0.035 | 0.035 | 0.04 | 0.04 | 0.049 | 0.042 | 0.050091 | 0.04
| Std Dev. |  | 0.0054772 | 0.0031984 | 0.0054772 | 7.6012e-18 | 0.0054772 | 0.0054772 | 0 | 0 | 0.004899 | 0.007975 | 0.00070786 | 0
| Rel.Std.Dev. |  | 12.17% | 7.655% | 9.959% | 1.52e-14% | 15.65% | 15.65% | 0% | 0% | 9.606% | 19.94% | 1.404% | 0%
| PDM3 |  | 3.141% | -4.232% | 26.06% | 14.6% | -19.78% | -19.78% | -8.319% | -8.319% | 16.89% | -8.319% | 15.53% | -8.319%
| Z-Score (Absolute) |  | 0.54 | 0.11 | 1.89 | 1.22 | 0.81 | 0.81 | 0.13 | 0.13 | 1.38 | 0.11 | 1.27 | 0.13
| Analytical results for Cl in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Cl | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.23 | [11 empty] | ppm
|  |  | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 11
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co3O4 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Co3O4 | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.23 | [11 empty] | ppm
|  |  | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 12
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 18
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 466 (Certified Value 197 ppm) | [13 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.18 | 2.19 | 2.23 | 2.25 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 100 | 180 | 322 | 100 | 200 | 224.6927 | 300 | 200 | 200 |  |  |  | 1
| 1 | 2 | < 100 | 190 | 307 | 200 | 200 | 231.5103 | 300 | 200 | 200 |  |  |  | 0
| 1 | 3 | < 100 | 160 | 395 | 100 | 200 | 848.2024 | 200 | 200 | 200 |  |  |  | 16
| 1 | 4 | < 100 | 150 | 322 | 100 | 200 | 223.8767 | 300 | 200 | 200 |  |  |  | 197.0764
| 1 | 5 | < 100 | 160 | 322 | 100 | 200 | 230.589 | 300 | 200 | 200 |  |  |  | 149
| 1 | 6 | 100 | 170 | 336 | 100 | 200 | 228.6723 | 200 | 200 | 200
| Mean |  | 100 | 168.3333 | 334 | 116.6667 | 200 | 331.2572 | 266.6667 | 200 | 200
| Median |  | 100 | 165 | 322 | 100 | 200 | 229.6307 | 300 | 200 | 200
| Std Dev. |  | 0 | 14.7196 | 31.2602 | 40.8248 | 0 | 253.269 | 51.6398 | 0 | 0
| Rel.Std.Dev. |  | 0% | 8.744% | 9.359% | 34.99% | 0% | 76.46% | 19.36% | 0% | 0%
| PDM3 |  | -49.26% | -14.58% | 69.48% | -40.8% | 1.483% | 68.09% | 35.31% | 1.483% | 1.483%
| Z-Score (Absolute) |  | 1.35 | 0.32 | 1.35 | 0.84 | 0 | 1.33 | 0.67 | 0 | 0
| Analytical results for Cs2O in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Cs2O | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.23 | [11 empty] | ppm
|  |  | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 13
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 19
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for CuO in OREAS 466 (Indicative Value 47 ppm) | [13 empty] | Indicative
CuO | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.07 | 2.18 | 2.23 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 1
1
| 1 | 1 | < 75 | 50 | 55.7035 | < 100 | [8 empty] | 1
| 1 | 2 | < 75 | 50 | 55.7694 | < 100 | [8 empty] | 14
| 1 | 3 | < 75 | 50 | 66.7028 | < 100 | [8 empty] | 16
| 1 | 4 | < 75 | 50 | 38.7079 | < 100 | [8 empty] | 46.9995
| 1 | 5 | < 75 | < 50 | 53.0711 | < 100 | [8 empty] | 20
| 1 | 6 | < 75 | < 50 | 44.0397 | < 100
| Mean |  |  | 50 | 52.3324
| Median |  |  | 50 | 54.3873
| Std Dev. |  |  | 0 | 9.8451
| Rel.Std.Dev. |  |  | 0% | 18.81%
| PDM3 |  |  | 6.384% | 11.35%
| Z-Score (Absolute) |  | 1.05 | 0.53 | 0.82 | 0.53
| Analytical results for Dy2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Dy2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 19.212 | < 100 |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 19.178 | < 100 |  |  |  | 0
| 1 | 3 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 19.557 | < 100 |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 19.431 | < 100 |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 18.696 | < 100 |  |  |  | 150
| 1 | 6 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 18.914 | < 100
| Mean |  | 2.11: 100; 2.24: 19.1647
| Median |  | 2.11: 100; 2.24: 19.195
| Std Dev. |  | 2.11: 0; 2.24: 0.3193
| Rel.Std.Dev. |  | 2.11: 0%; 2.24: 1.666%
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.25
| Analytical results for Er2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Er2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 15.918 | < 100 |  |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 15.529 | < 100 |  |  |  |  | 0
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 15.655 | < 100 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 15.574 | < 100 |  |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 15.003 | < 100 |  |  |  |  | 151
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 15.048 | < 100
| Mean |  | 2.24: 15.4545
| Median |  | 2.24: 15.5515
| Std Dev. |  | 2.24: 0.35888
| Rel.Std.Dev. |  | 2.24: 2.322%
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.25
| Analytical results for Eu2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Eu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 12 |  |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 12 |  |  |  |  | 0
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 12 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 12 |  |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 12 |  |  |  |  | 152
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 12
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.24
| Analytical results for F in OREAS 466 (Indicative Value < 0.5 wt.%) | [13 empty] | Indicative
F | Round | Replicate | 2.07: Lab |  |  |  |  | 1
| No | No | 2.07: 2.23 |  |  |  |  | wt.%
|  |  | 2.07: BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 2.07: < 0.5 |  |  |  |  | 1
| 1 | 2 | 2.07: < 0.5 |  |  |  |  | 15
| 1 | 3 | 2.07: < 0.5 |  |  |  |  | 16
| 1 | 4 | 2.07: < 0.5 |  |  |  |  | < 0.5
| 1 | 5 | 2.07: < 0.5 |  |  |  |  | 21
| 1 | 6 | 2.07: < 0.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07
| Analytical results for Fe2O3 in OREAS 466 (Certified Value 7.29 wt.%) | [13 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 7.38 | 6.88 | 7.717 | 10.4425 | 7.34 | 7.42 | 7.35 | 7.1128 | 7.2389 | 7.05 | 7.25 | 7.16 | 1
| 1 | 2 | 7.36 | 6.98 | 7.722 | 10.5269 | 7.38 | 7.31 | 7.39 | 7.2414 | 7.2469 | 7.07 | 7.3 | 7.34 | 0
| 1 | 3 | 7.38 | 7.06 | 7.653 | 10.454 | 7.36 | 7.41 | 7.37 | 7.2128 | 7.248 | 7.07 | 7.29 | 7.32 | 16
| 1 | 4 | 7.34 | 7.1 | 7.666 | 10.4683 | 7.36 | 7.27 | 7.37 | 7.22 | 7.2988 | 7.07 | 7.3 | 7.2 | 7.2873
| 1 | 5 | 7.42 | 6.99 | 7.648 | 10.5011 | 7.34 | 7.36 | 7.41 | 7.2414 | 7.3126 | 7.11 | 7.35 | 7.24 | 153
| 1 | 6 | 7.37 | 6.94 | 7.644 | 10.4325 | 7.4 | 7.34 | 7.41 | 7.2772 | 7.2719 | 7.14 | 7.26 | 7.23
| Mean |  | 7.375 | 6.9917 | 7.675 | 10.4709 | 7.3633 | 7.3517 | 7.3833 | 7.2176 | 7.2695 | 7.085 | 7.2917 | 7.2483
| Median |  | 7.375 | 6.985 | 7.6595 | 10.4611 | 7.36 | 7.35 | 7.38 | 7.2307 | 7.2599 | 7.07 | 7.295 | 7.235
| Std Dev. |  | 0.026646 | 0.079603 | 0.035293 | 0.036391 | 0.023381 | 0.057764 | 0.024221 | 0.05603 | 0.030424 | 0.033317 | 0.035449 | 0.069402
| Rel.Std.Dev. |  | 0.3613% | 1.139% | 0.4598% | 0.3475% | 0.3175% | 0.7857% | 0.3281% | 0.7763% | 0.4185% | 0.4702% | 0.4862% | 0.9575%
| PDM3 |  | 1.204% | -4.056% | 5.321% | 43.69% | 1.044% | 0.8836% | 1.318% | -0.9561% | -0.2435% | -2.776% | 0.06028% | -0.5344%
| Z-Score (Absolute) |  | 0.53 | 3.3 | 3.53 | 31.46 | 0.42 | 0.3 | 0.62 | 1.04 | 0.52 | 2.36 | 0.3 | 0.73
| Analytical results for Ga2O3 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Ga2O3 | Round | Replicate | 2.07: Lab |  |  |  |  | 1
| No | No | 2.07: 2.23 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 100 |  |  |  |  | 1
| 1 | 2 | 2.07: < 100 |  |  |  |  | 16
| 1 | 3 | 2.07: < 100 |  |  |  |  | 16
| 1 | 4 | 2.07: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 100 |  |  |  |  | 22
| 1 | 6 | 2.07: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07
| Analytical results for Gd2O3 in OREAS 466 (Indicative Value 110 ppm) | [13 empty] | Indicative
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 132.318 | < 100 |  |  | 1
| 1 | 2 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 132.353 | < 100 |  |  | 17
| 1 | 3 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 132.18 | < 100 |  |  | 16
| 1 | 4 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 129.852 | < 100 |  |  | 109.8796
| 1 | 5 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 125.576 | < 100 |  |  | 23
| 1 | 6 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 125.553 | < 100
| Mean |  | 2.07: 100; 2.12: 100 | 129.6387
| Median |  | 2.07: 100; 2.12: 100 | 131.016
| Std Dev. |  | 2.07: 0; 2.12: 0 | 3.2939
| Rel.Std.Dev. | [9 empty] | 2.541% | bar: 2.07; 2.12
| PDM3 |  | 2.07: -8.991%; 2.12: -8.991% | 17.98%
| Z-Score (Absolute) | [9 empty] | NA | NA | bar: 2.07 to 2.24
| Analytical results for HfO2 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
HfO2 | Round | Replicate | 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab |  |  |  |  | 1
| No | No | 2.07: 2.05; 2.08: 2.19; 2.10: 2.23; 2.11: 2.24 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF; 2.10: BF*XRF; 2.11: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 120; 2.08: < 100; 2.10: < 100; 2.11: < 12 |  |  |  |  | 1
| 1 | 2 | 2.07: < 120; 2.08: < 100; 2.10: < 100; 2.11: < 12 |  |  |  |  | 9
| 1 | 3 | 2.07: < 120; 2.08: < 100; 2.10: < 100; 2.11: < 12 |  |  |  |  | 16
| 1 | 4 | 2.07: < 120; 2.08: < 100; 2.10: < 100; 2.11: < 12 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 120; 2.08: < 100; 2.10: < 100; 2.11: < 12 |  |  |  |  | 15
| 1 | 6 | 2.07: < 120; 2.08: < 100; 2.10: < 100; 2.11: < 12
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.11
| Analytical results for Ho2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Ho2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.24 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 0
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 154
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.23
| Analytical results for K2O in OREAS 466 (Certified Value 2.87 wt.%) | [13 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.88 | 2.7 | 2.994 | 3.4705 | 2.92 | 2.86 | 2.91 | 2.765 | 2.7876 | > 1 | 2.85 | 2.94 | 1
| 1 | 2 | 2.87 | 2.69 | 2.998 | 3.497 | 2.89 | 2.84 | 2.91 | 2.795 | 2.7969 | > 1 | 2.86 | 2.92 | 0
| 1 | 3 | 2.86 | 2.72 | 2.986 | 3.4994 | 2.91 | 2.86 | 2.89 | 2.8 | 2.7806 | > 1 | 2.84 | 2.9 | 16
| 1 | 4 | 2.87 | 2.78 | 2.989 | 3.4777 | 2.9 | 2.82 | 2.9 | 2.815 | 2.778 | > 1 | 2.86 | 2.88 | 2.8724
| 1 | 5 | 2.88 | 2.69 | 2.988 | 3.5066 | 2.9 | 2.83 | 2.9 | 2.795 | 2.8057 | > 1 | 2.86 | 2.87 | 155
| 1 | 6 | 2.88 | 2.68 | 2.987 | 3.503 | 2.92 | 2.84 | 2.9 | 2.81 | 2.808 | > 1 | 2.84 | 2.87
| Mean |  | 2.8733 | 2.71 | 2.9903 | 3.4924 | 2.9067 | 2.8417 | 2.9017 | 2.7967 | 2.7928 |  | 2.8517 | 2.8967
| Median |  | 2.875 | 2.695 | 2.9885 | 3.4982 | 2.905 | 2.84 | 2.9 | 2.7975 | 2.7922 |  | 2.855 | 2.89
| Std Dev. |  | 0.008165 | 0.036878 | 0.0046762 | 0.014696 | 0.012111 | 0.016021 | 0.0075277 | 0.017512 | 0.012723 |  | 0.0098319 | 0.028752
| Rel.Std.Dev. |  | 0.2842% | 1.361% | 0.1564% | 0.4208% | 0.4166% | 0.5638% | 0.2594% | 0.6262% | 0.4556% |  | 0.3448% | 0.9926%
| PDM3 |  | 0.03302% | -5.653% | 4.106% | 21.58% | 1.193% | -1.069% | 1.019% | -2.636% | -2.771% |  | -0.7213% | 0.8454%
| Z-Score (Absolute) |  | 0 | 3.3 | 2.37 | 12.52 | 0.67 | 0.64 | 0.57 | 1.55 | 1.63 | NA | 0.44 | 0.47
| Analytical results for La2O3 in OREAS 466 (Certified Value 0.03 wt.%) | [13 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.03 | 0.0293 | 0.03 | 0.03 | 0.02 | 0.01 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03144 | 0.02 | 1
| 1 | 2 | 0.03 | 0.0293 | 0.03 | 0.03 | 0.03 | 0.01 | 0.02 | 0.03 | 0.03 | 0.03 | 0.030608 | 0.03 | 0
| 1 | 3 | 0.03 | 0.0305 | 0.03 | 0.03 | 0.02 | 0.01 | 0.03 | 0.03 | 0.03 | 0.03 | 0.030113 | 0.03 | 16
| 1 | 4 | 0.03 | 0.0281 | 0.03 | 0.03 | 0.03 | 0.01 | 0.03 | 0.03 | 0.03 | 0.03 | 0.030132 | 0.03 | 0.030001
| 1 | 5 | 0.03 | 0.0305 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.03 | 0.03 | 0.03 | 0.030122 | 0.03 | 156
| 1 | 6 | 0.03 | 0.0293 | 0.03 | 0.03 | 0.02 | 0.01 | 0.03 | 0.03 | 0.03 | 0.03 | 0.030544 | 0.03
| Mean |  | 0.03 | 0.0295 | 0.03 | 0.03 | 0.023333 | 0.011667 | 0.028333 | 0.03 | 0.03 | 0.03 | 0.030493 | 0.028333
| Median |  | 0.03 | 0.0293 | 0.03 | 0.03 | 0.02 | 0.01 | 0.03 | 0.03 | 0.03 | 0.03 | 0.030338 | 0.03
| Std Dev. |  | 0 | 0.00090333 | 0 | 0 | 0.005164 | 0.0040825 | 0.0040825 | 0 | 0 | 0 | 0.00051498 | 0.0040825
| Rel.Std.Dev. |  | 0% | 3.062% | 0% | 0% | 22.13% | 34.99% | 14.41% | 0% | 0% | 0% | 1.689% | 14.41%
| PDM3 |  | -0.002886% | -1.67% | -0.002886% | -0.002886% | -22.22% | -61.11% | -5.558% | -0.002886% | -0.002886% | -0.002886% | 1.641% | -5.558%
| Z-Score (Absolute) |  | 0 | 1.35 | 0 | 0 | 18.56 | 51.03 | 4.64 | 0 | 0 | 0 | 1.37 | 4.64
| Analytical results for Lu2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Lu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.24 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 1
| 1 | 2 | 200 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 0
| 1 | 3 | 200 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 16
| 1 | 4 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | < 100
| 1 | 5 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 157
| 1 | 6 | 200 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11
| Mean |  | 2.07: 150
| Median |  | 2.07: 150
| Std Dev. |  | 2.07: 54.7723
| Rel.Std.Dev. |  | 2.07: 36.51%
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.23
| Analytical results for MgO in OREAS 466 (Certified Value 0.824 wt.%) | [13 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.84 | 0.72 | 0.828 | 1.3315 | 0.87 | 0.8 | 0.83 | 0.8 | 0.82247 | 0.81 | 0.82 | 0.82 | 1
| 1 | 2 | 0.83 | 0.72 | 0.828 | 1.3398 | 0.86 | 0.81 | 0.84 | 0.815 | 0.82865 | 0.8 | 0.82 | 0.79 | 0
| 1 | 3 | 0.83 | 0.73 | 0.826 | 1.3514 | 0.88 | 0.81 | 0.84 | 0.81 | 0.82532 | 0.81 | 0.82 | 0.78 | 16
| 1 | 4 | 0.83 | 0.73 | 0.826 | 1.3746 | 0.87 | 0.81 | 0.85 | 0.805 | 0.82471 | 0.8 | 0.82 | 0.75 | 0.82357
| 1 | 5 | 0.83 | 0.71 | 0.825 | 1.3614 | 0.88 | 0.8 | 0.84 | 0.8 | 0.83433 | 0.82 | 0.83 | 0.79 | 158
| 1 | 6 | 0.82 | 0.7 | 0.825 | 1.3846 | 0.87 | 0.8 | 0.84 | 0.825 | 0.82992 | 0.81 | 0.82 | 0.8
| Mean |  | 0.83 | 0.71833 | 0.82633 | 1.3572 | 0.87167 | 0.805 | 0.84 | 0.80917 | 0.82757 | 0.80833 | 0.82167 | 0.78833
| Median |  | 0.83 | 0.72 | 0.826 | 1.3564 | 0.87 | 0.805 | 0.84 | 0.8075 | 0.82699 | 0.81 | 0.82 | 0.79
| Std Dev. |  | 0.0063246 | 0.01169 | 0.0013663 | 0.020337 | 0.0075277 | 0.0054772 | 0.0063246 | 0.009704 | 0.0042811 | 0.0075277 | 0.0040825 | 0.023166
| Rel.Std.Dev. |  | 0.762% | 1.627% | 0.1653% | 1.498% | 0.8636% | 0.6804% | 0.7529% | 1.199% | 0.5173% | 0.9313% | 0.4969% | 2.939%
| PDM3 |  | 0.7803% | -12.78% | 0.3351% | 64.8% | 5.84% | -2.255% | 1.995% | -1.749% | 0.4849% | -1.85% | -0.2315% | -4.279%
| Z-Score (Absolute) |  | 0.26 | 4.5 | 0.1 | 22.71 | 2.03 | 0.81 | 0.68 | 0.63 | 0.15 | 0.67 | 0.1 | 1.52
| Analytical results for MnO in OREAS 466 (Certified Value 0.011 wt.%) | [13 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.19 | 2.23 | 2.25 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.02 | < 0.01 | 0.013 | 0.011 | 0.02 | 0.01 | 0.01 | 0.01 | 0.009 | 0.01 |  |  | 1
| 1 | 2 | 0.02 | < 0.01 | 0.015 | 0.013 | 0.02 | 0.01 | < 0.01 | < 0.01 | 0.019 | 0.01 |  |  | 0
| 1 | 3 | 0.02 | < 0.01 | 0.017 | 0.014 | 0.02 | 0.01 | 0.01 | 0.01 | 0.009 | < 0.01 |  |  | 16
| 1 | 4 | 0.02 | < 0.01 | 0.018 | 0.012 | 0.02 | 0.01 | 0.01 | 0.01 | 0.009 | 0.01 |  |  | 0.011486
| 1 | 5 | 0.02 | < 0.01 | 0.017 | 0.011 | 0.02 | 0.01 | 0.01 | < 0.01 | 0.009 | 0.01 |  |  | 159
| 1 | 6 | 0.02 | < 0.01 | 0.015 | 0.012 | 0.02 | 0.01 | 0.01 | 0.01 | 0.019 | 0.01
| Mean |  | 0.02 |  | 0.015833 | 0.012167 | 0.02 | 0.01 | 0.01 | 0.01 | 0.012333 | 0.01
| Median |  | 0.02 |  | 0.016 | 0.012 | 0.02 | 0.01 | 0.01 | 0.01 | 0.009 | 0.01
| Std Dev. |  | 0 |  | 0.0018348 | 0.001169 | 0 | 0 | 0 | 0 | 0.005164 | 0
| Rel.Std.Dev. |  | 0% |  | 11.59% | 9.609% | 0% | 0% | 0% | 0% | 41.87% | 0%
| PDM3 |  | 74.13% |  | 37.85% | 5.927% | 74.13% | -12.94% | -12.94% | -12.94% | 7.378% | -12.94%
| Z-Score (Absolute) |  | 2.58 | 1.76 | 1.37 | 0.31 | 2.58 | 0.31 | 0.55 | 0.79 | 0.38 | 0.55
| Analytical results for MoO3 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
MoO3 | Round | Replicate | 2.07: Lab |  |  |  |  | 1
| No | No | 2.07: 2.23 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 100 |  |  |  |  | 1
| 1 | 2 | 2.07: < 100 |  |  |  |  | 10
| 1 | 3 | 2.07: < 100 |  |  |  |  | 16
| 1 | 4 | 2.07: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 100 |  |  |  |  | 16
| 1 | 6 | 2.07: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07
| Analytical results for Na2O in OREAS 466 (Certified Value 0.236 wt.%) | [13 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.09 | 2.16 | 2.18 | 2.23 | 2.25 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.28 | 0.21 | 0.191 | 0.3289 | 0.14 | 0.25 | 0.211 | 0.26676 | 0.25 | 0.2 |  |  | 1
| 1 | 2 | 0.28 | 0.23 | 0.193 | 0.3451 | 0.13 | 0.24 | 0.202 | 0.25973 | 0.25 | < 0.1 |  |  | 7
| 1 | 3 | 0.29 | 0.23 | 0.171 | 0.3249 | 0.13 | 0.25 | 0.2245 | 0.26135 | 0.24 | < 0.1 |  |  | 16
| 1 | 4 | 0.28 | 0.23 | 0.178 | 0.3491 | 0.13 | 0.24 | 0.217 | 0.26167 | 0.24 | < 0.1 |  |  | 0.23613
| 1 | 5 | 0.28 | 0.23 | 0.184 | 0.3451 | 0.13 | 0.25 | 0.2225 | 0.27106 | 0.24 | < 0.1 |  |  | 160
| 1 | 6 | 0.28 | 0.2 | 0.182 | 0.3383 | 0.12 | 0.24 | 0.217 | 0.26369 | 0.23 | < 0.1
| Mean |  | 0.28167 | 0.22167 | 0.18317 | 0.33857 | 0.13 | 0.245 | 0.21567 | 0.26404 | 0.24167 | 0.2
| Median |  | 0.28 | 0.23 | 0.183 | 0.3417 | 0.13 | 0.245 | 0.217 | 0.26268 | 0.24 | 0.2
| Std Dev. |  | 0.0040825 | 0.013292 | 0.0081833 | 0.0097633 | 0.0063246 | 0.0054772 | 0.0082077 | 0.0042011 | 0.0075277
| Rel.Std.Dev. |  | 1.449% | 5.996% | 4.468% | 2.884% | 4.865% | 2.236% | 3.806% | 1.591% | 3.115%
| PDM3 |  | 19.29% | -6.123% | -22.43% | 43.38% | -44.94% | 3.758% | -8.664% | 11.82% | 2.347% | -15.3%
| Z-Score (Absolute) |  | 1.04 | 0.52 | 1.52 | 2.51 | 2.9 | 0.09 | 0.67 | 0.58 | 0 | NA
Results from laboratory 25 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb2O5 in OREAS 466 (Certified Value 147 ppm) | [13 empty] | Certified
Nb2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 140 | 114 | 140 | 200 | 186 | 114 | 143 | 186 | 140 | 200 | 161.389 | < 500 | 1
| 1 | 2 | 140 | 114 | 130 | 100 | 200 | 114 | 129 | 215 | < 140 | 300 | 155.266 | < 500 | 0
| 1 | 3 | 290 | 100 | 160 | 200 | 186 | 114 | 129 | 172 | 140 | 200 | 154.909 | < 500 | 16
| 1 | 4 | 140 | 100 | 70 | 100 | 200 | 114 | 129 | 186 | 140 | 200 | 149.816 | < 500 | 147.3543
| 1 | 5 | 290 | 143 | 170 | 100 | 200 | 129 | 143 | 200 | < 140 | 200 | 149.201 | < 500 | 161
| 1 | 6 | 140 | 114 | 120 | 100 | 200 | 114 | 143 | 200 | 140 | 200 | 145.267 | < 500
| Mean |  | 190 | 114.1667 | 131.6667 | 133.3333 | 195.3333 | 116.5 | 136 | 193.1667 | 140 | 216.6667 | 152.6413
| Median |  | 140 | 114 | 135 | 100 | 200 | 114 | 136 | 193 | 140 | 200 | 152.3625
| Std Dev. |  | 77.4597 | 15.7024 | 35.4495 | 51.6398 | 7.2296 | 6.1237 | 7.6681 | 14.9722 | 0 | 40.8248 | 5.7053
| Rel.Std.Dev. |  | 40.77% | 13.75% | 26.92% | 38.73% | 3.701% | 5.256% | 5.638% | 7.751% | 0% | 18.84% | 3.738%
| PDM3 |  | 28.94% | -22.52% | -10.65% | -9.515% | 32.56% | -20.94% | -7.705% | 31.09% | -4.991% | 47.04% | 3.588%
| Z-Score (Absolute) |  | 1.09 | 0.7 | 0.3 | NA | 1.21 | 0.65 | 0.2 | 1.15 | 0.59 | NA | 0.2 | 2.49
Results from laboratories 8 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for Nd2O3 in OREAS 466 (Indicative Value 0.017 wt.%) | [13 empty] | Indicative
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.01 | 0.0187 | 0.03 | 0.01 | < 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | 0.02 | 0.019143 | < 0.01 | 1
| 1 | 2 | 0.01 | 0.0198 | 0.03 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | 0.02 | 0.018871 | < 0.01 | 11
| 1 | 3 | 0.01 | 0.0152 | 0.03 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | < 0.02 | 0.02 | 0.019684 | < 0.01 | 16
| 1 | 4 | 0.01 | 0.0117 | 0.03 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | < 0.02 | < 0.01 | 0.020087 | < 0.01 | 0.016942
| 1 | 5 | 0.01 | 0.014 | 0.02 | 0.01 | < 0.01 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.020011 | < 0.01 | 17
| 1 | 6 | 0.01 | 0.0152 | 0.03 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | < 0.02 | 0.01 | 0.019926 | < 0.01
| Mean |  | 0.01 | 0.015767 | 0.028333 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.026667 | 0.016 | 0.01962
| Median |  | 0.01 | 0.0152 | 0.03 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | 0.02 | 0.019805
| Std Dev. |  | 0 | 0.0030058 | 0.0040825 | 0 | 0 | 0 | 0 | 0 | 0.0057735 | 0.0054772 | 0.00050124
| Rel.Std.Dev. |  | 0% | 19.06% | 14.41% | 0% | 0% | 0% | 0% | 0% | 21.65% | 34.23% | 2.555%
| PDM3 |  | -40.98% | -6.94% | 67.23% | -40.98% | -40.98% | -40.98% | 18.05% | 18.05% | 57.4% | -5.562% | 15.81%
| Z-Score (Absolute) |  | 0.67 | 0.11 | 1.82 | 0.67 | 0.9 | 0.67 | 0.69 | 0.69 | 0.46 | 0.11 | 0.63 | 1.35
| Analytical results for NiO in OREAS 466 (Indicative Value 38.5 ppm) | [13 empty] | Indicative
NiO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.05 | 2.07 | 2.08 | 2.18 | 2.23 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 1
1
| 1 | 1 | < 76 | 60 | < 130 | 27.4485 | < 100 | < 100 | [6 empty] | 1
| 1 | 2 | < 76 | 60 | < 130 | 25.2355 | < 100 | < 100 | [6 empty] | 12
| 1 | 3 | < 76 | 70 | < 130 | 80.6458 | < 100 | < 100 | [6 empty] | 16
| 1 | 4 | < 76 | 40 | < 130 | 21.9871 | < 100 | < 100 | [6 empty] | 38.5363
| 1 | 5 | < 76 | 60 | < 130 | 23.1331 | < 100 | < 100 | [6 empty] | 18
| 1 | 6 | < 76 | 30 | < 130 | 20.8925 | < 100 | < 100
| Mean |  | 2.08: 53.3333; 2.11: 33.2237
| Median |  | 2.08: 60; 2.11: 24.1843
| Std Dev. |  | 2.08: 15.0555; 2.11: 23.3502
| Rel.Std.Dev. |  | 2.08: 28.23%; 2.11: 70.28%
| PDM3 |  | 2.08: 38.4%; 2.11: -13.79%
| Z-Score (Absolute) |  | 1.05 | 0.3 | 1.21 | 1.49 | 0 | 0
| Analytical results for P2O5 in OREAS 466 (Certified Value 0.126 wt.%) | [13 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.127 | 0.12 | 0.129 | 0.291 | 0.13 | 0.11 | 0.13 | 0.1175 | 0.12248 | 0.12 | 0.13 | 0.13 | 1
| 1 | 2 | 0.132 | 0.12 | 0.129 | 0.2887 | 0.12 | 0.11 | 0.13 | 0.121 | 0.12111 | 0.12 | 0.13 | 0.12 | 0
| 1 | 3 | 0.132 | 0.12 | 0.129 | 0.291 | 0.13 | 0.11 | 0.13 | 0.1215 | 0.12235 | 0.12 | 0.13 | 0.14 | 16
| 1 | 4 | 0.131 | 0.13 | 0.128 | 0.2887 | 0.14 | 0.11 | 0.13 | 0.1205 | 0.1217 | 0.12 | 0.13 | 0.13 | 0.12647
| 1 | 5 | 0.132 | 0.13 | 0.131 | 0.3025 | 0.13 | 0.11 | 0.13 | 0.12 | 0.12209 | 0.12 | 0.13 | 0.13 | 162
| 1 | 6 | 0.128 | 0.12 | 0.13 | 0.3002 | 0.13 | 0.11 | 0.13 | 0.117 | 0.12304 | 0.12 | 0.13 | 0.13
| Mean |  | 0.13033 | 0.12333 | 0.12933 | 0.29368 | 0.13 | 0.11 | 0.13 | 0.11958 | 0.12213 | 0.12 | 0.13 | 0.13
| Median |  | 0.1315 | 0.12 | 0.129 | 0.291 | 0.13 | 0.11 | 0.13 | 0.12025 | 0.12222 | 0.12 | 0.13 | 0.13
| Std Dev. |  | 0.0022509 | 0.005164 | 0.0010328 | 0.0060707 | 0.0063246 | 0 | 0 | 0.0018819 | 0.0006677 | 0 | 0 | 0.0063246
| Rel.Std.Dev. |  | 1.727% | 4.187% | 0.7986% | 2.067% | 4.865% | 0% | 0% | 1.574% | 0.5467% | 0% | 0% | 4.865%
| PDM3 |  | 3.054% | -2.481% | 2.263% | 132.2% | 2.79% | -13.02% | 2.79% | -5.446% | -3.434% | -5.117% | 2.79% | 2.79%
| Z-Score (Absolute) |  | 0.13 | 1.22 | 0.06 | 31.6 | 0.06 | 3.79 | 0.06 | 1.94 | 1.45 | 1.86 | 0.06 | 0.06
| Analytical results for PbO in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
PbO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.05 | 2.18 | 2.19 | 2.23 | 2.25 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 86 | 33.7207 | 100 | < 100 | < 97 | [7 empty] | 1
| 1 | 2 | < 86 | 34.4602 | < 100 | < 100 | < 97 | [7 empty] | 13
| 1 | 3 | < 86 | 32.742 | < 100 | < 100 | < 97 | [7 empty] | 16
| 1 | 4 | < 86 | 29.4545 | < 100 | < 100 | < 97 | [7 empty] | < 100
| 1 | 5 | < 86 | 33.3833 | < 100 | < 100 | < 97 | [7 empty] | 19
| 1 | 6 | < 86 | 29.3117 | 100 | < 100 | < 97
| Mean |  | 2.08: 32.1787; 2.10: 100
| Median |  | 2.08: 33.0627; 2.10: 100
| Std Dev. |  | 2.08: 2.2357; 2.10: 0
| Rel.Std.Dev. |  | 2.08: 6.948%; 2.10: 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 2.04 | 2.28 | 0.19 | 0
| Analytical results for Pr6O11 in OREAS 466 (Indicative Value 112 ppm) | [13 empty] | Indicative
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | 100 | < 100 | < 100 | 100 | 200 | 100 | < 200 | < 100 | 60.446 | < 100 |  |  | 1
| 1 | 2 | 100 | < 100 | < 100 | 100 | 200 | 100 | < 200 | < 100 | 59.178 | < 100 |  |  | 14
| 1 | 3 | 100 | < 100 | < 100 | 100 | 200 | 100 | < 200 | < 100 | 58.706 | < 100 |  |  | 16
| 1 | 4 | 100 | < 100 | < 100 | 100 | 200 | 100 | < 200 | < 100 | 59.081 | < 100 |  |  | 111.7896
| 1 | 5 | 100 | < 100 | < 100 | 100 | 200 | 100 | < 200 | < 100 | 58.803 | < 100 |  |  | 20
| 1 | 6 | 100 | < 100 | < 100 | 100 | 200 | 100 | < 200 | < 100 | 58.972 | < 100
| Mean |  | 2.07: 100; 2.11: 100; 2.12: 200; 2.13: 100 | 59.1977
| Median |  | 2.07: 100; 2.11: 100; 2.12: 200; 2.13: 100 | 59.0265
| Std Dev. |  | 2.07: 0; 2.11: 0; 2.12: 0; 2.13: 0 | 0.63573
| Rel.Std.Dev. | [9 empty] | 1.074% | bar: 2.07; 2.11 to 2.13
| PDM3 |  | 2.07: -10.55%; 2.11: -10.55%; 2.12: 78.91%; 2.13: -10.55% | -47.05%
| Z-Score (Absolute) |  | 0.55 | 0.8 | 0.8 | 0.55 | 3.25 | 0.55 | 0.55 | 0.8 | 0.55 | 0.8
| Analytical results for Rb2O in OREAS 466 (Indicative Value 133 ppm) | [13 empty] | Indicative
Rb2O | Round | Replicate | 2.07: Lab; 2.08: Lab |  |  |  |  | 1
| No | No | 2.07: 2.07; 2.08: 2.23 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 100; 2.08: 200 |  |  |  |  | 1
| 1 | 2 | 2.07: 100; 2.08: 100 |  |  |  |  | 15
| 1 | 3 | 2.07: 100; 2.08: 200 |  |  |  |  | 16
| 1 | 4 | 2.07: < 100; 2.08: 200 |  |  |  |  | 133.3333
| 1 | 5 | 2.07: 100; 2.08: 200 |  |  |  |  | 21
| 1 | 6 | 2.07: 100; 2.08: 200
| Mean |  | 2.07: 100; 2.08: 183.3333
| Median |  | 2.07: 100; 2.08: 200
| Std Dev. |  | 2.07: 0; 2.08: 40.8248
| Rel.Std.Dev. |  | 2.07: 0%; 2.08: 22.27%
| PDM3 |  | 2.07: -25%; 2.08: 37.5%
| Z-Score (Absolute) |  | 2.07: 0.67; 2.08: 0.67
| Analytical results for Sb2O3 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Sb2O3 | Round | Replicate | 2.07: Lab |  |  |  |  | 1
| No | No | 2.07: 2.23 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 100 |  |  |  |  | 1
| 1 | 2 | 2.07: < 100 |  |  |  |  | 16
| 1 | 3 | 2.07: < 100 |  |  |  |  | 16
| 1 | 4 | 2.07: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 100 |  |  |  |  | 22
| 1 | 6 | 2.07: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07
| Analytical results for Sc2O3 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Sc2O3 | Round | Replicate | 2.07: Lab; 2.08: Lab |  |  |  |  | 1
| No | No | 2.07: 2.08; 2.08: 2.23 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 150; 2.08: < 100 |  |  |  |  | 1
| 1 | 2 | 2.07: < 150; 2.08: < 100 |  |  |  |  | 17
| 1 | 3 | 2.07: < 150; 2.08: < 100 |  |  |  |  | 16
| 1 | 4 | 2.07: < 150; 2.08: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 150; 2.08: < 100 |  |  |  |  | 23
| 1 | 6 | 2.07: < 150; 2.08: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.08
| Analytical results for SiO2 in OREAS 466 (Certified Value 66.4 wt.%) | [13 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 66.93 | 65.2 | 66.028 | 139.3625 | 66.54 | 66.6 | 67.12 | 63.085 | 65.7319 | 66.02 | 66.41 | 66.7 | 1
| 1 | 2 | 66.82 | 65.32 | 66.008 | 140.6802 | 66.61 | 66.65 | 67.17 | 64.76 | 65.8792 | 65.9 | 66.66 | 66.9 | 0
| 1 | 3 | 66.79 | 66.38 | 66.141 | 140.5283 | 66.62 | 66.7 | 67.15 | 63.99 | 65.7771 | 66.01 | 66.15 | 66.9 | 16
| 1 | 4 | 66.76 | 67.13 | 66.096 | 139.5892 | 66.46 | 66.59 | 67.19 | 64.455 | 65.8138 | 66.17 | 66.3 | 66 | 66.403
| 1 | 5 | 67.06 | 65.45 | 66.212 | 140.5839 | 66.48 | 66.56 | 67.12 | 64.49 | 66.2284 | 65.96 | 66.95 | 66.2 | 163
| 1 | 6 | 66.76 | 65.15 | 66.213 | 140.3294 | 67.04 | 66.59 | 67.19 | 64.8 | 65.8226 | 66.16 | 66.31 | 66.4
| Mean |  | 66.8533 | 65.7717 | 66.1163 | 140.1789 | 66.625 | 66.615 | 67.1567 | 64.2633 | 65.8755 | 66.0367 | 66.4633 | 66.5167
| Median |  | 66.805 | 65.385 | 66.1185 | 140.4289 | 66.575 | 66.595 | 67.16 | 64.4725 | 65.8182 | 66.015 | 66.36 | 66.55
| Std Dev. |  | 0.11928 | 0.80447 | 0.088383 | 0.56113 | 0.21352 | 0.050892 | 0.032042 | 0.64591 | 0.17971 | 0.1082 | 0.29214 | 0.37639
| Rel.Std.Dev. |  | 0.1784% | 1.223% | 0.1337% | 0.4003% | 0.3205% | 0.0764% | 0.04771% | 1.005% | 0.2728% | 0.1638% | 0.4396% | 0.5659%
| PDM3 |  | 0.6782% | -0.9508% | -0.4317% | 111.1% | 0.3343% | 0.3192% | 1.135% | -3.222% | -0.7944% | -0.5517% | 0.09083% | 0.1711%
| Z-Score (Absolute) |  | 0.59 | 1.17 | 0.61 | 120.17 | 0.22 | 0.2 | 1.09 | 3.63 | 1 | 0.74 | 0.04 | 0.04
| Analytical results for Sm2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 300 | < 100 | 24.734 | < 100 |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 300 | < 100 | 24.908 | < 100 |  |  | 0
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | 100 | < 300 | < 100 | 24.63 | < 100 |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 300 | < 100 | 24.34 | < 100 |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 300 | < 100 | 23.308 | < 100 |  |  | 164
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 300 | < 100 | 23.343 | < 100
| Mean |  | 2.13: 100 | 24.2105
| Median |  | 2.13: 100 | 24.485
| Std Dev. | [9 empty] | 0.70998
| Rel.Std.Dev. | [9 empty] | 2.933%
PDM3
| Z-Score (Absolute) | [9 empty] | NA | NA | bar: 2.07 to 2.24
| Analytical results for SnO2 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
SnO2 | Round | Replicate | 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab |  |  |  |  | 1
| No | No | 2.07: 2.08; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF; 2.10: BF*XRF; 2.11: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 100; 2.08: < 100; 2.10: < 100; 2.11: < 100 |  |  |  |  | 1
| 1 | 2 | 2.07: < 100; 2.08: < 100; 2.10: < 100; 2.11: < 100 |  |  |  |  | 0
| 1 | 3 | 2.07: < 100; 2.08: < 100; 2.10: < 100; 2.11: < 100 |  |  |  |  | 16
| 1 | 4 | 2.07: < 100; 2.08: < 100; 2.10: < 100; 2.11: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 100; 2.08: < 100; 2.10: < 100; 2.11: < 100 |  |  |  |  | 165
| 1 | 6 | 2.07: < 100; 2.08: 100; 2.10: < 100; 2.11: < 100
| Mean |  | 2.08: 100
| Median |  | 2.08: 100
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.11
| Analytical results for SO3 in OREAS 466 (Indicative Value 0.013 wt.%) | [13 empty] | Indicative
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.09 | 2.23 | [7 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 3
3
| 1 | 1 | 0.01 | < 0.01 | 0.0674 | 0.017 | 0.02 | [7 empty] | 1
| 1 | 2 | 0.007 | < 0.01 | 0.0325 | 0.017 | 0.02 | [7 empty] | 9
| 1 | 3 | 0.006 | < 0.01 | 0.035 | 0.012 | 0.02 | [7 empty] | 16
| 1 | 4 | 0.008 | 0.02 | 0.0499 | 0.012 | 0.02 | [7 empty] | 0.012625
| 1 | 5 | 0.01 | < 0.01 | 0.0549 | 0.015 | 0.02 | [7 empty] | 15
| 1 | 6 | 0.008 | 0.01 | 0.0325 | 0.011 | 0.02
| Mean |  | 0.0081667 | 0.015 | 0.045367 | 0.014 | 0.02
| Median |  | 0.008 | 0.015 | 0.04245 | 0.0135 | 0.02
| Std Dev. |  | 0.0016021 | 0.0070711 | 0.014391 | 0.0026833 | 0
| Rel.Std.Dev. |  | 19.62% | 47.14% | 31.72% | 19.17% | 0%
| PDM3 |  | -35.31% | 18.81% | 259.3% | 10.89% | 58.42%
| Z-Score (Absolute) |  | 0.67 | 0.66 | 3.63 | 0 | 0.69
| Analytical results for SrO in OREAS 466 (Certified Value 112 ppm) | [13 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.09 | 2.18 | 2.23 | 2.25 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 200 | < 120 | 120 | 130 | 120.6941 | 100 | 100 | [5 empty] | 1
| 1 | 2 | 200 | < 120 | 110 | 142 | 122.128 | < 100 | 100 | [5 empty] | 0
| 1 | 3 | 100 | < 120 | 130 | 130 | 125.261 | 100 | 100 | [5 empty] | 16
| 1 | 4 | 200 | < 120 | 110 | 189 | 122.8176 | < 100 | 100 | [5 empty] | 112.3179
| 1 | 5 | 200 | < 120 | 120 | 142 | 124.2175 | 100 | 100 | [5 empty] | 166
| 1 | 6 | 200 | < 120 | 120 | 142 | 121.3305 | 100 | 100
| Mean |  | 183.3333 |  | 118.3333 | 145.8333 | 122.7415 | 100 | 100
| Median |  | 200 |  | 120 | 142 | 122.4728 | 100 | 100
| Std Dev. |  | 40.8248 |  | 7.5277 | 21.9492 | 1.7373 | 0 | 0
| Rel.Std.Dev. |  | 22.27% |  | 6.361% | 15.05% | 1.415% | 0% | 0%
| PDM3 |  | 63.23% |  | 5.356% | 29.84% | 9.28% | -10.97% | -10.97%
| Z-Score (Absolute) |  | 1.59 | 1.45 | 0 | 0.67 | 0.11 | 0.86 | 0.45
| Analytical results for Ta2O5 in OREAS 466 (Indicative Value < 100 ppm) | [13 empty] | Indicative
Ta2O5 | Round | Replicate | 2.07: Lab; 2.08: Lab; 2.10: Lab |  |  |  |  | 1
| No | No | 2.07: 2.07; 2.08: 2.23; 2.10: 2.25 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF; 2.10: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: < 30; 2.08: < 100; 2.10: < 500 |  |  |  |  | 1
| 1 | 2 | 2.07: < 30; 2.08: < 100; 2.10: < 500 |  |  |  |  | 10
| 1 | 3 | 2.07: < 30; 2.08: < 100; 2.10: < 500 |  |  |  |  | 16
| 1 | 4 | 2.07: < 30; 2.08: < 100; 2.10: < 500 |  |  |  |  | < 100
| 1 | 5 | 2.07: < 30; 2.08: < 100; 2.10: < 500 |  |  |  |  | 16
| 1 | 6 | 2.07: < 30; 2.08: < 100; 2.10: < 500
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.10
| Analytical results for Tb4O7 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Tb4O7 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.24 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 62.409 | [5 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 61.48 | [5 empty] | 0
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 62.433 | [5 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 61.586 | [5 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60.468 | [5 empty] | 167
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 61.127
| Mean |  | 2.23: 61.5838
| Median |  | 2.23: 61.533
| Std Dev. |  | 2.23: 0.75717
| Rel.Std.Dev. |  | 2.23: 1.229%
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.23
| Analytical results for ThO2 in OREAS 466 (Indicative Value 56 ppm) | [13 empty] | Indicative
ThO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 57 | 120 | < 100 | 68 | < 57 | 57 | < 57 | < 100 | < 100 | 41.943 | < 100 |  | 1
| 1 | 2 | 57 | 60 | < 100 | 80 | < 57 | 57 | < 57 | < 100 | < 100 | 42.25 | < 100 |  | 11
| 1 | 3 | < 57 | 100 | < 100 | 68 | < 57 | < 57 | < 57 | < 100 | < 100 | 40.885 | < 100 |  | 16
| 1 | 4 | < 57 | 100 | < 100 | 68 | < 57 | < 57 | < 57 | < 100 | < 100 | 42.227 | < 100 |  | 55.9239
| 1 | 5 | < 57 | 110 | < 100 | 68 | < 57 | 57 | < 57 | < 100 | < 100 | 41.795 | < 100 |  | 17
| 1 | 6 | < 57 | 50 | < 100 | 80 | < 57 | 57 | < 57 | < 100 | < 100 | 41.055 | < 100
| Mean |  | 2.07: 57; 2.08: 90; 2.11: 72; 2.13: 57 |  | 41.6925
| Median |  | 2.07: 57; 2.08: 100; 2.11: 68; 2.13: 57 |  | 41.869
| Std Dev. |  | 2.08: 28.2843; 2.11: 6.1968; 2.13: 0 |  | 0.58791
| Rel.Std.Dev. |  | 2.08: 31.43%; 2.11: 8.607%; 2.13: 0% |  | 1.41%
| PDM3 |  | 2.07: 1.924%; 2.08: 60.93%; 2.11: 28.75%; 2.13: 1.924% |  | -25.45%
| Z-Score (Absolute) |  | 1.36 | 3.25 | 0 | 1.79 | 1.75 | 0.21 | 1.75 | 0 | 0 | 0.67 | 0
| Analytical results for TiO2 in OREAS 466 (Certified Value 0.986 wt.%) | [13 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.16 | 2.18 | 2.19 | 2.23 | 2.25 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.99 | 0.95 | 1.059 | 1.6014 | 1 | 1.02 | 1 | 0.97 | 0.98521 | 0.98 | 0.98 | 0.98 | 1
| 1 | 2 | 0.99 | 0.97 | 1.056 | 1.6164 | 1 | 0.99 | 1 | 0.98 | 0.98248 | 0.97 | 0.98 | 1 | 0
| 1 | 3 | 0.99 | 0.96 | 1.053 | 1.603 | 1.01 | 1.01 | 0.99 | 0.985 | 0.98156 | 0.97 | 0.98 | 0.98 | 16
| 1 | 4 | 0.99 | 0.97 | 1.051 | 1.6047 | 1 | 0.99 | 1 | 0.98 | 0.98783 | 0.97 | 0.97 | 0.98 | 0.98577
| 1 | 5 | 0.99 | 0.96 | 1.046 | 1.6097 | 1 | 1.01 | 1 | 0.985 | 1.0006 | 0.97 | 0.99 | 0.99 | 168
| 1 | 6 | 0.99 | 0.95 | 1.049 | 1.6131 | 1.01 | 1.01 | 1 | 0.965 | 0.99362 | 0.97 | 0.98 | 0.97
| Mean |  | 0.99 | 0.96 | 1.0523 | 1.608 | 1.0033 | 1.005 | 0.99833 | 0.9775 | 0.98855 | 0.97167 | 0.98 | 0.98333
| Median |  | 0.99 | 0.96 | 1.052 | 1.6072 | 1 | 1.01 | 1 | 0.98 | 0.98652 | 0.97 | 0.98 | 0.98
| Std Dev. |  | 0 | 0.0089443 | 0.0047188 | 0.0059816 | 0.005164 | 0.012247 | 0.0040825 | 0.0082158 | 0.0073329 | 0.0040825 | 0.0063246 | 0.010328
| Rel.Std.Dev. |  | 0% | 0.9317% | 0.4484% | 0.372% | 0.5147% | 1.219% | 0.4089% | 0.8405% | 0.7418% | 0.4202% | 0.6454% | 1.05%
| PDM3 |  | 0.4289% | -2.614% | 6.752% | 63.13% | 1.781% | 1.951% | 1.274% | -0.8392% | 0.2823% | -1.431% | -0.5855% | -0.2474%
| Z-Score (Absolute) |  | 0.04 | 1.53 | 3.29 | 32.3 | 0.73 | 0.82 | 0.47 | 0.61 | 0.04 | 0.92 | 0.48 | 0.31
| Analytical results for Tm2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Tm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.24 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 0
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11 | [5 empty] | 169
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 11
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.23
| Analytical results for TOT_XRF in OREAS 466 (Indicative Value 99.34 wt.%) | [13 empty] | Indicative
TOT_XRF | Round | Replicate | 2.07: Lab; 2.08: Lab; 2.10: Lab |  |  |  |  | 1
| No | No | 2.07: 2.10; 2.08: 2.11; 2.10: 2.12 |  |  |  |  | wt.%
|  |  | 2.07: BF*XRF; 2.08: BF*XRF; 2.10: BF*XRF |  |  |  |  | 2
2
| 1 | 1 | 2.07: 98.98; 2.08: 98.16; 2.10: 100.45 |  |  |  |  | 1
| 1 | 2 | 2.07: 98.75; 2.08: 98.96; 2.10: 99.81 |  |  |  |  | 12
| 1 | 3 | 2.07: 99.28; 2.08: 99.86; 2.10: 100.15 |  |  |  |  | 16
| 1 | 4 | 2.07: 98.66; 2.08: 98.55; 2.10: 100.35 |  |  |  |  | 99.335
| 1 | 5 | 2.07: 98.59; 2.08: 100.25; 2.10: 99.7 |  |  |  |  | 18
| 1 | 6 | 2.07: 98.75; 2.08: 99.42; 2.10: 99.36
| Mean |  | 2.07: 98.835; 2.08: 99.2; 2.10: 99.97
| Median |  | 2.07: 98.75; 2.08: 99.19; 2.10: 99.98
| Std Dev. |  | 2.07: 0.25462; 2.08: 0.7935; 2.10: 0.419
| Rel.Std.Dev. |  | 2.07: 0.2576%; 2.08: 0.7999%; 2.10: 0.4191%
| PDM3 |  | 2.07: -0.5033%; 2.08: -0.1359%; 2.10: 0.6393%
| Z-Score (Absolute) |  | 2.07: 0.67; 2.08: 0; 2.10: 1.42
| Analytical results for U3O8 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
U3O8 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 59 | < 50 | < 100 | 24 | < 24 | < 24 | < 24 | < 100 | < 100 | 4.21 | < 100 |  | 1
| 1 | 2 | < 59 | < 50 | < 100 | < 24 | < 24 | < 24 | 24 | < 100 | < 100 | 4.245 | < 100 |  | 0
| 1 | 3 | < 59 | < 50 | < 100 | 35 | < 24 | < 24 | 24 | < 100 | < 100 | 4.304 | < 100 |  | 16
| 1 | 4 | < 59 | < 50 | < 100 | < 24 | < 24 | < 24 | < 24 | < 100 | < 100 | 4.222 | < 100 |  | < 100
| 1 | 5 | < 59 | < 50 | < 100 | < 24 | < 24 | 24 | < 24 | < 100 | < 100 | 4.257 | < 100 |  | 170
| 1 | 6 | < 59 | < 50 | < 100 | 24 | < 24 | 24 | 24 | < 100 | < 100 | 4.222 | < 100
| Mean |  | 2.11: 27.6667; 2.13: 24; 2.23: 24 |  | 4.2433
| Median |  | 2.11: 24; 2.13: 24; 2.23: 24 |  | 4.2335
| Std Dev. |  | 2.11: 6.3509; 2.13: 0; 2.23: 0 |  | 0.034326
| Rel.Std.Dev. |  | 2.11: 22.95%; 2.13: 0%; 2.23: 0% |  | 0.8089%
PDM3
| Z-Score (Absolute) |  | 0.23 | 0 | 1.28 | 0.27 | 0.67 | 0.47 | 0.37 | 1.28 | 1.28 | 1.06 | 1.28
| Analytical results for V2O5 in OREAS 466 (Indicative Value 244 ppm) | [13 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.05 | 2.19 | 2.23 | 2.25 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 100 | 428 | 200 | 300 | 200 | [7 empty] | 1
| 1 | 2 | 200 | 411 | 300 | 300 | 200 | [7 empty] | 13
| 1 | 3 | 100 | 428 | 300 | 200 | 200 | [7 empty] | 16
| 1 | 4 | < 100 | 446 | 200 | 300 | 200 | [7 empty] | 244.4444
| 1 | 5 | 100 | 464 | 200 | 300 | 200 | [7 empty] | 19
| 1 | 6 | 200 | 393 | 300 | 200 | 300
| Mean |  | 140 | 428.3333 | 250 | 266.6667 | 216.6667
| Median |  | 100 | 428 | 250 | 300 | 200
| Std Dev. |  | 54.7723 | 25.0333 | 54.7723 | 51.6398 | 40.8248
| Rel.Std.Dev. |  | 39.12% | 5.844% | 21.91% | 19.36% | 18.84%
| PDM3 |  | -42.73% | 75.23% | 2.273% | 9.091% | -11.36%
| Z-Score (Absolute) |  | 2.53 | 3.61 | 0 | 0.34 | 0.67
| Analytical results for WO3 in OREAS 466 (Indicative Value 121 ppm) | [13 empty] | Indicative
WO3 | Round | Replicate | 2.07: Lab; 2.08: Lab |  |  |  |  | 1
| No | No | 2.07: 2.23; 2.08: 2.25 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: 100; 2.08: 100 |  |  |  |  | 1
| 1 | 2 | 2.07: 100; 2.08: 200 |  |  |  |  | 14
| 1 | 3 | 2.07: 100; 2.08: 100 |  |  |  |  | 16
| 1 | 4 | 2.07: 100; 2.08: < 100 |  |  |  |  | 120.8333
| 1 | 5 | 2.07: 100; 2.08: 200 |  |  |  |  | 20
| 1 | 6 | 2.07: 100; 2.08: 200
| Mean |  | 2.07: 100; 2.08: 160
| Median |  | 2.07: 100; 2.08: 200
| Std Dev. |  | 2.07: 0; 2.08: 54.7723
| Rel.Std.Dev. |  | 2.07: 0%; 2.08: 34.23%
| PDM3 |  | 2.07: -17.24%; 2.08: 32.41%
| Z-Score (Absolute) |  | 2.07: 0.67; 2.08: 0.67
| Analytical results for Y2O3 in OREAS 466 (Indicative Value 83 ppm) | [13 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.19 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | 100 | < 130 | < 130 | 130 | < 130 | < 100 | < 100 | 31.621 | < 100 |  |  | 1
| 1 | 2 | < 100 | 100 | < 130 | < 130 | 130 | 130 | < 100 | < 100 | 30.947 | < 100 |  |  | 15
| 1 | 3 | < 100 | 100 | < 130 | < 130 | 130 | < 130 | < 100 | < 100 | 31.849 | < 100 |  |  | 16
| 1 | 4 | < 100 | 100 | < 130 | < 130 | 130 | < 130 | < 100 | < 100 | 32.078 | < 100 |  |  | 83.1403
| 1 | 5 | < 100 | 100 | < 130 | < 130 | 130 | < 130 | < 100 | < 100 | 31.354 | < 100 |  |  | 21
| 1 | 6 | < 100 | 100 | < 130 | < 130 | 130 | < 130 | < 100 | < 100 | 31.113 | < 100
| Mean |  | 2.08: 100; 2.12: 130; 2.13: 130 | 31.4937
| Median |  | 2.08: 100; 2.12: 130; 2.13: 130 | 31.4875
| Std Dev. |  | 2.08: 0; 2.12: 0 | 0.43524
| Rel.Std.Dev. | [9 empty] | 1.382% | bar: 2.08; 2.12
| PDM3 |  | 2.08: 20.28%; 2.12: 56.36%; 2.13: 56.36% | -62.12%
| Z-Score (Absolute) |  | 0.67 | 4.32 | 0.67 | 0.67 | 7.02 | 1.73 | 0.67 | 0.67 | 2.52 | 0.67
| Analytical results for Yb2O3 in OREAS 466 (Certified Value < 100 ppm) | [13 empty] | Certified
Yb2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 11 | < 100 |  |  |  |  | 1
| 1 | 2 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 11 | < 100 |  |  |  |  | 0
| 1 | 3 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 11 | < 100 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 11 | < 100 |  |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 11 | < 100 |  |  |  |  | 171
| 1 | 6 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 11 | < 100
| Mean |  | 2.10: 100
| Median |  | 2.10: 100
| Std Dev. |  | 2.10: 0
Rel.Std.Dev. | bar: 2.10
PDM3
Z-Score (Absolute) | bar: 2.07 to 2.25
| Analytical results for ZnO in OREAS 466 (Indicative Value 59 ppm) | [13 empty] | Indicative
ZnO | Round | Replicate | 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab |  |  |  |  | 1
| No | No | 2.07: 2.05; 2.08: 2.18; 2.10: 2.23; 2.11: 2.25 |  |  |  |  | ppm
|  |  | 2.07: BF*XRF; 2.08: BF*XRF; 2.10: BF*XRF; 2.11: BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2.07: 75; 2.08: 43.4781; 2.10: < 100; 2.11: < 100 |  |  |  |  | 1
| 1 | 2 | 2.07: 62; 2.08: 44.1207; 2.10: < 100; 2.11: < 100 |  |  |  |  | 16
| 1 | 3 | 2.07: 75; 2.08: 44.0689; 2.10: < 100; 2.11: < 100 |  |  |  |  | 16
| 1 | 4 | 2.07: 87; 2.08: 43.9936; 2.10: < 100; 2.11: < 100 |  |  |  |  | 59.204
| 1 | 5 | 2.07: 87; 2.08: 43.5201; 2.10: < 100; 2.11: < 100 |  |  |  |  | 22
| 1 | 6 | 2.07: 62; 2.08: 43.1384; 2.10: < 100; 2.11: < 100
| Mean |  | 2.07: 74.6667; 2.08: 43.72
| Median |  | 2.07: 75; 2.08: 43.7568
| Std Dev. |  | 2.07: 11.1833; 2.08: 0.39849
| Rel.Std.Dev. |  | 2.07: 14.98%; 2.08: 0.9115%
| PDM3 |  | 2.07: 26.12%; 2.08: -26.15%
| Z-Score (Absolute) |  | 2.07: 5.3; 2.08: 1.35; 2.10: 0; 2.11: 0
| Analytical results for ZrO2 in OREAS 466 (Certified Value 363 ppm) | [13 empty] | Certified
ZrO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.05 | 2.07 | 2.08 | 2.18 | 2.19 | 2.23 | 2.25 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 300 | 554 | 330 | 200 | 364.6988 | 400 | 400 | 400 |  |  |  |  | 1
| 1 | 2 | 310 | 500 | 330 | 200 | 364.6626 | 400 | 400 | 400 |  |  |  |  | 0
| 1 | 3 | 310 | 554 | 350 | 200 | 359.5701 | 300 | 400 | 400 |  |  |  |  | 16
| 1 | 4 | 310 | 621 | 320 | 200 | 365.6768 | 400 | 400 | 400 |  |  |  |  | 363.2181
| 1 | 5 | 330 | 513 | 310 | 200 | 368.0918 | 400 | 400 | 400 |  |  |  |  | 172
| 1 | 6 | 310 | 540 | 380 | 200 | 363.1499 | 400 | 400 | 300
| Mean |  | 311.6667 | 547 | 336.6667 | 200 | 364.3083 | 383.3333 | 400 | 383.3333
| Median |  | 310 | 547 | 330 | 200 | 364.6807 | 400 | 400 | 400
| Std Dev. |  | 9.8319 | 42.3981 | 25.0333 | 0 | 2.8375 | 40.8248 | 0 | 40.8248
| Rel.Std.Dev. |  | 3.155% | 7.751% | 7.436% | 0% | 0.7789% | 10.65% | 0% | 10.65%
| PDM3 |  | -14.19% | 50.6% | -7.31% | -44.94% | 0.3002% | 5.538% | 10.13% | 5.538%
| Z-Score (Absolute) |  | 1.32 | 3.69 | 0.79 | 3.7 | 0.2 | 0.2 | 0.56 | 0.2
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C649: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C667: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C703: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C721: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C739: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C757: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C775: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C793: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C811: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C829: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C847: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C865: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C883: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C901: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C919: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C937: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C955: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C973: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 466 (Indicative Value 0.642 wt.%) | [18 empty] | Indicative
H2O- | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 2.18 | 2.23 | [15 empty] | wt.%
|  |  | BF*XRF | LOI*TGA | [15 empty] | 3
3
| 1 | 1 | 0.7246 | 0.67 | [15 empty] | 1
| 1 | 2 | 0.6843 | 0.59 | [15 empty] | 19
| 1 | 3 | 0.6117 | 0.6 | [15 empty] | 16
| 1 | 4 | 0.6747 | 0.61 | [15 empty] | 0.64245
| 1 | 5 | 0.658 | 0.62 | [15 empty] | 25
| 1 | 6 | 0.6661 | 0.6
| Mean |  | 0.6699 | 0.615
| Median |  | 0.6704 | 0.605
| Std Dev. |  | 0.036785 | 0.02881
| Rel.Std.Dev. |  | 5.491% | 4.685%
| PDM3 |  | 4.273% | -4.273%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for LOI1000 in OREAS 466 (Certified Value 4 wt.%) | [18 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.16 | 2.18 | 2.19 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | BF*XRF | BF*XRF | BF*XRF | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | 2
3
| 1 | 1 | 3.73 | 4.17 | 4.4 | 5.76 | 4.42 | 3.81 | 3.8 | 3.65 | 4.17 | 3.74 | 3.64 | 3.81 | 4.3861 | 4.3521 | 3.8 | 4.25 | 3.65 | 1
| 1 | 2 | 3.73 | 4.19 | 4.4 | 4.87 | 4.39 | 3.87 | 3.7 | 3.6 | 4.5 | 3.76 | 3.69 | 3.76 | 4.3346 | 4.4867 | 3.79 | 4.29 | 3.71 | 0
| 1 | 3 | 3.75 | 4.18 | 4.32 | 5.17 | 4.44 | 3.81 | 3.8 | 3.62 | 4.45 | 3.74 | 3.6 | 3.8 | 4.2661 | 4.3799 | 3.8 | 4.3 | 3.67 | 16
| 1 | 4 | 3.79 | 4.23 | 4.32 | 5.39 | 4.42 | 3.84 | 3.8 | 3.68 | 4.47 | 3.75 | 3.62 | 3.75 | 4.3262 | 4.3792 | 3.79 | 4.35 | 3.84 | 4.0004
| 1 | 5 | 3.72 | 4.17 | 4.32 | 5.08 | 4.44 | 3.86 | 3.8 | 3.64 | 4.3 | 3.75 | 3.7 | 3.78 | 4.3199 | 4.4015 | 3.77 | 4.21 | 3.64 | 174
| 1 | 6 | 3.73 | 4.26 | 4.32 | 5.28 | 4.31 | 3.8 | 3.7 | 3.66 | 4.5 | 3.76 | 3.66 | 3.81 | 4.3212 | 4.4142 | 3.78 | 4.2 | 3.73
| Mean |  | 3.7417 | 4.2 | 4.3467 | 5.2583 | 4.4033 | 3.8317 | 3.7667 | 3.6417 | 4.3983 | 3.75 | 3.6517 | 3.785 | 4.3257 | 4.4023 | 3.7883 | 4.2667 | 3.7067
| Median |  | 3.73 | 4.185 | 4.32 | 5.225 | 4.42 | 3.825 | 3.8 | 3.645 | 4.46 | 3.75 | 3.65 | 3.79 | 4.3237 | 4.3907 | 3.79 | 4.27 | 3.69
| Std Dev. |  | 0.025626 | 0.036878 | 0.041312 | 0.30328 | 0.049261 | 0.029269 | 0.05164 | 0.028577 | 0.13438 | 0.0089443 | 0.0392 | 0.025884 | 0.038298 | 0.046529 | 0.01169 | 0.057504 | 0.073937
| Rel.Std.Dev. |  | 0.6849% | 0.8781% | 0.9504% | 5.768% | 1.119% | 0.7639% | 1.371% | 0.7847% | 3.055% | 0.2385% | 1.073% | 0.6839% | 0.8854% | 1.057% | 0.3086% | 1.348% | 1.995%
| PDM3 |  | -6.467% | 4.99% | 8.656% | 31.45% | 10.07% | -4.218% | -5.843% | -8.967% | 9.948% | -6.259% | -8.717% | -5.384% | 8.131% | 10.05% | -5.301% | 6.656% | -7.342%
| Z-Score (Absolute) |  | 0.32 | 1.31 | 1.83 | 5.06 | 2.03 | 0 | 0.23 | 0.67 | 2.01 | 0.29 | 0.64 | 0.17 | 1.75 | 2.03 | 0.15 | 1.54 | 0.44
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.