## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 143
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 12.3832 | 12.206 | 12.5603 | 12.2468 | 12.5196
As, Arsenic (wt.%) | 0.11372 | 0.10207 | 0.12536 | IND | IND
BaO, Barium oxide (ppm) | 737.7848 | 642.7308 | 832.8387 | 715.5091 | 760.0604
CaO, Calcium oxide (wt.%) | 9.7931 | 9.6032 | 9.983 | 9.6791 | 9.9071
Co, Cobalt (ppm) | < 100 | IND | IND | IND | IND
Cu, Copper (ppm) | 319.2899 | 276.5466 | 362.0332 | 290.4426 | 348.1373
Fe2O3, Iron(III) oxide (wt.%) | 6.9844 | 6.8536 | 7.1152 | 6.8935 | 7.0752
HfO2, Hafnium dioxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 2.8376 | 2.7836 | 2.8917 | 2.8134 | 2.8619
MgO, Magnesium oxide (wt.%) | 2.0511 | 2.0123 | 2.0899 | 2.018 | 2.0841
MnO, Manganese oxide (wt.%) | 0.058288 | 0.050108 | 0.066468 | 0.055929 | 0.060648
Mo, Molybdenum (ppm) | < 50 | IND | IND | IND | IND
Na2O, Sodium oxide (wt.%) | 1.2131 | 1.1143 | 1.3119 | 1.192 | 1.2341
P2O5, Phosphorus(V) oxide (wt.%) | 0.12658 | 0.11861 | 0.13455 | 0.12329 | 0.12988
SiO2, Silicon dioxide (wt.%) | 54.1658 | 53.4621 | 54.8696 | 53.6478 | 54.6838
Sn, Tin (wt.%) | 0.25212 | 0.24076 | 0.26347 | 0.24315 | 0.26109
SO3, Sulphur trioxide (wt.%) | 3.0759 | 2.9934 | 3.1583 | 3.0223 | 3.1295
SrO, Strontium oxide (ppm) | 120.1151 | 89.4104 | 150.8198 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.57564 | 0.55949 | 0.59179 | 0.5616 | 0.58968
V2O5, Vanadium(V) oxide (ppm) | 135.2579 | 94.1928 | 176.323 | IND | IND
Zn, Zinc (ppm) | 108.4297 | 90.6007 | 126.2587 | IND | IND
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 6.6584 | 6.4846 | 6.8323 | 6.5162 | 6.8006
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.515 | 6.2354 | 6.7946 | 6.3765 | 6.6534
As, Arsenic (wt.%) | 0.11533 | 0.11079 | 0.11987 | 0.11303 | 0.11763
Ba, Barium (ppm) | 649.4503 | 624.7448 | 674.1558 | 636.8 | 662.1007
Be, Beryllium (ppm) | 3.1441 | 2.9193 | 3.369 | IND | IND
Ca, Calcium (wt.%) | 7.0494 | 6.6583 | 7.4405 | 6.8824 | 7.2164
Ce, Cerium (ppm) | 66.9466 | 61.5305 | 72.3626 | 63.8546 | 70.0386
Co, Cobalt (ppm) | 13.6125 | 11.8379 | 15.387 | 13.0008 | 14.2242
Cr, Chromium (ppm) | 74.3025 | 60.1247 | 88.4802 | IND | IND
Cs, Caesium (ppm) | 9.3816 | 8.587 | 10.1762 | 8.7559 | 10.0073
Cu, Copper (ppm) | 332.5379 | 317.3117 | 347.764 | 324.9804 | 340.0953
Dy, Dysprosium (ppm) | 5.0102 | 4.5317 | 5.4887 | 4.7122 | 5.3082
Er, Erbium (ppm) | 2.7537 | 2.2697 | 3.2377 | 2.5277 | 2.9797
Eu, Europium (ppm) | 1.2193 | 1.0394 | 1.3992 | IND | IND
Fe, Iron (wt.%) | 4.8545 | 4.686 | 5.023 | 4.7723 | 4.9367
Ga, Gallium (ppm) | 17.8989 | 15.9949 | 19.8029 | 17.2487 | 18.5492
Gd, Gadolinium (ppm) | 5.3284 | 4.5073 | 6.1495 | 5.0509 | 5.6059
Ho, Holmium (ppm) | 0.95534 | 0.74096 | 1.1697 | IND | IND
K, Potassium (wt.%) | 2.3788 | 2.2642 | 2.4934 | 2.3235 | 2.4341
La, Lanthanum (ppm) | 32.5111 | 29.1042 | 35.918 | 31.2434 | 33.7788
Li, Lithium (ppm) | 59.0658 | 54.8209 | 63.3107 | 56.6519 | 61.4797
Mg, Magnesium (wt.%) | 1.1994 | 1.1441 | 1.2546 | 1.1738 | 1.2249
Mn, Manganese (wt.%) | 0.04561 | 0.043928 | 0.047291 | 0.044975 | 0.046244
Nd, Neodymium (ppm) | 30.4628 | 28.4039 | 32.5216 | 29.3641 | 31.5615
Ni, Nickel (ppm) | 36.4081 | 24.9202 | 47.8959 | IND | IND
P, Phosphorus (wt.%) | 0.05748 | 0.051201 | 0.06376 | IND | IND
Pr, Praseodymium (ppm) | 7.8221 | 6.909 | 8.7352 | 7.4415 | 8.2028
Rb, Rubidium (ppm) | 151.2258 | 145.6869 | 156.7647 | 147.1097 | 155.3419
S, Sulphur (wt.%) | 1.2478 | 1.2041 | 1.2915 | 1.226 | 1.2696
Si, Silicon (wt.%) | 25.8014 | 24.4751 | 27.1277 | 25.0997 | 26.5032
Sm, Samarium (ppm) | 6.1289 | 5.4366 | 6.8211 | 5.7062 | 6.5515
Sn, Tin (wt.%) | 0.23755 | 0.22275 | 0.25235 | 0.23174 | 0.24336
Sr, Strontium (ppm) | 131.2679 | 123.9012 | 138.6346 | 127.695 | 134.8408
Ta, Tantalum (ppm) | 0.99606 | 0.60976 | 1.3824 | IND | IND
Tb, Terbium (ppm) | 0.80155 | 0.66907 | 0.93402 | IND | IND
Th, Thorium (ppm) | 13.0093 | 11.8547 | 14.164 | 12.6589 | 13.3597
Ti, Titanium (wt.%) | 0.34649 | 0.33115 | 0.36182 | 0.33462 | 0.35835
Tm, Thulium (ppm) | 0.37374 | 0.27265 | 0.47483 | IND | IND
U, Uranium (ppm) | 3.0795 | 2.4402 | 3.7188 | 2.8676 | 3.2914
V, Vanadium (ppm) | 69.9663 | 58.4428 | 81.4899 | 67.2564 | 72.6763
Y, Yttrium (ppm) | 25.6143 | 22.6487 | 28.5798 | 24.6862 | 26.5424
Yb, Ytterbium (ppm) | 2.42 | 1.8257 | 3.0143 | 2.09 | 2.7499
Zn, Zinc (ppm) | 107.5281 | 91.5226 | 123.5336 | 104.8023 | 110.2539
4-Acid Digestion
Ag, Silver (ppm) | 0.49727 | 0.44584 | 0.54871 | 0.4644 | 0.53015
Al, Aluminium (wt.%) | 6.3685 | 6.1974 | 6.5395 | 6.2573 | 6.4796
As, Arsenic (wt.%) | 0.116 | 0.11167 | 0.12033 | 0.11308 | 0.11891
Ba, Barium (ppm) | 655.7836 | 638.7652 | 672.8019 | 638.0267 | 673.5404
Be, Beryllium (ppm) | 3.0464 | 2.8178 | 3.2749 | 2.872 | 3.2207
Bi, Bismuth (ppm) | 23.3702 | 21.6526 | 25.0877 | 21.6757 | 25.0646
Ca, Calcium (wt.%) | 6.8532 | 6.6437 | 7.0628 | 6.7357 | 6.9708
Ce, Cerium (ppm) | 67.0281 | 62.3436 | 71.7127 | 65.7094 | 68.3469
Co, Cobalt (ppm) | 13.0289 | 12.1361 | 13.9217 | 12.3809 | 13.6769
Cr, Chromium (ppm) | 67.9856 | 62.2315 | 73.7397 | 64.2305 | 71.7407
Cs, Caesium (ppm) | 9.3066 | 8.8384 | 9.7747 | 8.9681 | 9.6451
Cu, Copper (ppm) | 327.7761 | 320.1597 | 335.3925 | 321.6364 | 333.9158
Dy, Dysprosium (ppm) | 3.0797 | 2.7533 | 3.4062 | 2.9353 | 3.2241
Er, Erbium (ppm) | 1.4582 | 1.2389 | 1.6774 | 1.3602 | 1.5561
Eu, Europium (ppm) | 1.1124 | 1.0143 | 1.2104 | 1.0113 | 1.2135
Fe, Iron (wt.%) | 4.7907 | 4.6886 | 4.8927 | 4.6843 | 4.897
Ga, Gallium (ppm) | 17.8847 | 16.9492 | 18.8202 | 17.3672 | 18.4021
Gd, Gadolinium (ppm) | 4.626 | 4.1247 | 5.1274 | 4.4073 | 4.8448
Hf, Hafnium (ppm) | 2.6627 | 2.4296 | 2.8958 | 2.489 | 2.8363
Ho, Holmium (ppm) | 0.52131 | 0.45247 | 0.59016 | 0.46926 | 0.57336
In, Indium (ppm) | 0.87151 | 0.65754 | 1.0855 | 0.76877 | 0.97424
K, Potassium (wt.%) | 2.3554 | 2.2995 | 2.4112 | 2.3008 | 2.4099
La, Lanthanum (ppm) | 33.2382 | 30.9927 | 35.4836 | 32.3839 | 34.0925
Li, Lithium (ppm) | 58.9647 | 56.4686 | 61.4607 | 57.5301 | 60.3993
Lu, Lutetium (ppm) | 0.19187 | 0.16667 | 0.21706 | IND | IND
Mg, Magnesium (wt.%) | 1.1509 | 1.1094 | 1.1923 | 1.1262 | 1.1755
Mn, Manganese (wt.%) | 0.044617 | 0.042984 | 0.046249 | 0.043654 | 0.04558
Mo, Molybdenum (ppm) | 1.3203 | 1.172 | 1.4687 | 1.1882 | 1.4525
Na, Sodium (wt.%) | 0.92567 | 0.89864 | 0.95269 | 0.90427 | 0.94707
Nb, Niobium (ppm) | 11.2766 | 10.4168 | 12.1363 | 10.6056 | 11.9475
Nd, Neodymium (ppm) | 28.28 | 25.3621 | 31.1979 | 26.2576 | 30.3024
Ni, Nickel (ppm) | 33.8525 | 32.0394 | 35.6656 | 32.5943 | 35.1107
P, Phosphorus (wt.%) | 0.057834 | 0.055572 | 0.060096 | 0.056661 | 0.059007
Pb, Lead (ppm) | 27.0692 | 24.8959 | 29.2424 | 25.5046 | 28.6338
Pr, Praseodymium (ppm) | 8.0388 | 7.2402 | 8.8375 | 7.5333 | 8.5444
Rb, Rubidium (ppm) | 146.0925 | 138.88 | 153.305 | 142.5554 | 149.6296
S, Sulphur (wt.%) | 1.2245 | 1.1882 | 1.2609 | 1.1978 | 1.2513
Sb, Antimony (ppm) | 1.5409 | 1.3659 | 1.7159 | 1.3678 | 1.714
Sc, Scandium (ppm) | 10.108 | 9.553 | 10.6629 | 9.639 | 10.5769
Sm, Samarium (ppm) | 5.6545 | 5.1575 | 6.1514 | 5.2362 | 6.0727
Sr, Strontium (ppm) | 128.0104 | 123.4667 | 132.554 | 124.9325 | 131.0882
Ta, Tantalum (ppm) | 0.95618 | 0.85788 | 1.0545 | 0.87482 | 1.0375
Tb, Terbium (ppm) | 0.60025 | 0.54721 | 0.65329 | 0.58039 | 0.62011
Th, Thorium (ppm) | 12.9408 | 12.1798 | 13.7018 | 12.4514 | 13.4303
Ti, Titanium (wt.%) | 0.32478 | 0.3098 | 0.33976 | 0.31767 | 0.3319
Tl, Thallium (ppm) | 0.78852 | 0.75497 | 0.82206 | 0.75944 | 0.81759
Tm, Thulium (ppm) | 0.17964 | 0.15635 | 0.20292 | IND | IND
U, Uranium (ppm) | 2.8044 | 2.4754 | 3.1333 | 2.5849 | 3.0238
V, Vanadium (ppm) | 71.734 | 69.5125 | 73.9556 | 69.8635 | 73.6046
W, Tungsten (ppm) | 21.3278 | 19.4147 | 23.2408 | 19.7446 | 22.9109
Y, Yttrium (ppm) | 13.7663 | 12.6361 | 14.8965 | 13.1654 | 14.3671
Yb, Ytterbium (ppm) | 1.2672 | 1.1003 | 1.4341 | 1.1851 | 1.3493
Zn, Zinc (ppm) | 99.8651 | 95.1702 | 104.5599 | 95.7937 | 103.9364
Zr, Zirconium (ppm) | 86.0146 | 80.0096 | 92.0196 | 83.1547 | 88.8745
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 143
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Bi | ppm | 80.9167 | Hg | ppm | < 100 | Sc | ppm | < 100
Cd | ppm | < 10 | In | ppm | < 100 | Se | ppm | < 100
Ce | ppm | < 80 | La | ppm | < 90 | Ta | ppm | < 100
Cl | ppm | < 50 | Nb | ppm | < 50 | Te | ppm | < 100
Cr2O3 | ppm | 137.4072 | NiO | ppm | < 64 | Tl | ppm | < 100
Cs | ppm | < 100 | Pb | ppm | 84.5486 | W | ppm | 15.4167
Ga | ppm | < 100 | Rb | ppm | 108.25 | Y | ppm | 47.6415
Ge | ppm | < 100 | Sb | ppm | < 50 | ZrO2 | ppm | 270.9479
Thermogravimetry
H2O- | wt.% | 0.295
Borate / Peroxide Fusion ICP
Ag | ppm | < 0.5 | In | ppm | 4.9225 | Sb | ppm | 2.4779
B | ppm | 663.9 | Lu | ppm | 0.38064 | Sc | ppm | 9.8187
Bi | ppm | 23.65 | Mo | ppm | 2.9444 | Se | ppm | < 8
Cd | ppm | 29.3003 | Na | wt.% | 0.51157 | Te | ppm | < 1
Ge | ppm | 1.7816 | Nb | ppm | 12.0228 | Tl | ppm | 0.86111
Hf | ppm | 4.9282 | Pb | ppm | 40.2667 | W | ppm | 26.719
Hg | ppm | < 5 | Re | ppm | < 0.1 | Zr | ppm | 191.6798
4-Acid Digestion
B | ppm | 7.2133 | Hg | ppm | 0.16 | Sn | wt.% | 0.013489
Cd | ppm | 0.46287 | Re | ppm | < 0.002 | Te | ppm | 0.031944
Ge | ppm | 0.1481 | Se | ppm | 1.7011
Infrared Combustion
S | wt.% | 1.23
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 143
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 12.3832 | 0.21292 | 11.9573 | 12.809 | 11.7444 | 13.022 | 1.719% | 3.439% | 5.158% | 11.764 | 13.0023
As, wt.% | 0.11372 | 0.0090206 | 0.095675 | 0.13176 | 0.086655 | 0.14078 | 7.933% | 15.87% | 23.8% | 0.10803 | 0.1194
BaO, ppm | 737.7848 | 66.4881 | 604.8085 | 870.761 | 538.3204 | 937.2491 | 9.012% | 18.02% | 27.04% | 700.8955 | 774.674
CaO, wt.% | 9.7931 | 0.16275 | 9.4676 | 10.1186 | 9.3048 | 10.2813 | 1.662% | 3.324% | 4.986% | 9.3034 | 10.2827
Co, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cu, ppm | 319.2899 | 25.5603 | 268.1694 | 370.4104 | 242.6091 | 395.9707 | 8.005% | 16.01% | 24.02% | 303.3254 | 335.2544
Fe2O3, wt.% | 6.9844 | 0.19148 | 6.6014 | 7.3674 | 6.41 | 7.5588 | 2.742% | 5.483% | 8.225% | 6.6352 | 7.3336
HfO2, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 2.8376 | 0.038393 | 2.7608 | 2.9144 | 2.7225 | 2.9528 | 1.353% | 2.706% | 4.059% | 2.6958 | 2.9795
MgO, wt.% | 2.0511 | 0.039283 | 1.9725 | 2.1296 | 1.9332 | 2.1689 | 1.915% | 3.831% | 5.746% | 1.9485 | 2.1536
MnO, wt.% | 0.058288 | 0.008097 | 0.042094 | 0.074482 | 0.033997 | 0.082579 | 13.89% | 27.78% | 41.67% | 0.055374 | 0.061203
Mo, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Na2O, wt.% | 1.2131 | 0.086624 | 1.0398 | 1.3863 | 0.95319 | 1.4729 | 7.141% | 14.28% | 21.42% | 1.1524 | 1.2737
P2O5, wt.% | 0.12658 | 0.0058126 | 0.11496 | 0.13821 | 0.10915 | 0.14402 | 4.592% | 9.184% | 13.78% | 0.12025 | 0.13291
SiO2, wt.% | 54.1658 | 0.75597 | 52.6539 | 55.6778 | 51.8979 | 56.4337 | 1.396% | 2.791% | 4.187% | 51.4575 | 56.8741
Sn, wt.% | 0.25212 | 0.0066375 | 0.23884 | 0.26539 | 0.23221 | 0.27203 | 2.633% | 5.265% | 7.898% | 0.23951 | 0.26472
SO3, wt.% | 3.0759 | 0.08144 | 2.913 | 3.2387 | 2.8315 | 3.3202 | 2.648% | 5.295% | 7.943% | 2.9221 | 3.2297
SrO, ppm | 120.1151 | 27.232 | 65.6511 | 174.5791 | 38.419 | 201.8111 | 22.67% | 45.34% | 68.01% | 114.1093 | 126.1208
TiO2, wt.% | 0.57564 | 0.010223 | 0.55519 | 0.59609 | 0.54497 | 0.60631 | 1.776% | 3.552% | 5.328% | 0.54686 | 0.60442
V2O5, ppm | 135.2579 | 25.0758 | 85.1063 | 185.4095 | 60.0306 | 210.4853 | 18.54% | 37.08% | 55.62% | 128.495 | 142.0208
Zn, ppm | 108.4297 | 20.7405 | 66.9487 | 149.9107 | 46.2082 | 170.6512 | 19.13% | 38.26% | 57.38% | 103.0082 | 113.8512
Thermogravimetry
LOI1000, wt.% | 6.6584 | 0.17689 | 6.3046 | 7.0122 | 6.1277 | 7.1891 | 2.657% | 5.313% | 7.97% | 6.3255 | 6.9913
Borate / Peroxide Fusion ICP
Al, wt.% | 6.515 | 0.18663 | 6.1417 | 6.8882 | 5.9551 | 7.0749 | 2.865% | 5.729% | 8.594% | 6.1892 | 6.8407
As, wt.% | 0.11533 | 0.0056604 | 0.10401 | 0.12665 | 0.098346 | 0.13231 | 4.908% | 9.816% | 14.72% | 0.10956 | 0.12109
Ba, ppm | 649.4503 | 20.0238 | 609.4026 | 689.498 | 589.3788 | 709.5218 | 3.083% | 6.166% | 9.25% | 616.9778 | 681.9228
Be, ppm | 3.1441 | 0.34071 | 2.4627 | 3.8256 | 2.122 | 4.1663 | 10.84% | 21.67% | 32.51% | 2.9869 | 3.3014
Ca, wt.% | 7.0494 | 0.24054 | 6.5683 | 7.5305 | 6.3278 | 7.771 | 3.412% | 6.824% | 10.24% | 6.6969 | 7.4019
Ce, ppm | 66.9466 | 3.7465 | 59.4536 | 74.4396 | 55.7071 | 78.1861 | 5.596% | 11.19% | 16.79% | 63.5992 | 70.2939
Co, ppm | 13.6125 | 1.1615 | 11.2895 | 15.9354 | 10.128 | 17.0969 | 8.533% | 17.07% | 25.6% | 12.9318 | 14.2931
Cr, ppm | 74.3025 | 8.2598 | 57.7829 | 90.8221 | 49.523 | 99.0819 | 11.12% | 22.23% | 33.35% | 70.5874 | 78.0176
Cs, ppm | 9.3816 | 0.42733 | 8.5269 | 10.2362 | 8.0996 | 10.6636 | 4.555% | 9.11% | 13.66% | 8.9125 | 9.8507
Cu, ppm | 332.5379 | 11.0679 | 310.402 | 354.6737 | 299.3341 | 365.7417 | 3.328% | 6.657% | 9.985% | 315.911 | 349.1648
Dy, ppm | 5.0102 | 0.24585 | 4.5185 | 5.5019 | 4.2727 | 5.7477 | 4.907% | 9.814% | 14.72% | 4.7597 | 5.2607
Er, ppm | 2.7537 | 0.35459 | 2.0445 | 3.4629 | 1.69 | 3.8175 | 12.88% | 25.75% | 38.63% | 2.616 | 2.8914
Eu, ppm | 1.2193 | 0.12256 | 0.97418 | 1.4644 | 0.85161 | 1.587 | 10.05% | 20.1% | 30.16% | 1.1583 | 1.2803
Fe, wt.% | 4.8545 | 0.10849 | 4.6375 | 5.0715 | 4.529 | 5.18 | 2.235% | 4.47% | 6.705% | 4.6118 | 5.0972
Ga, ppm | 17.8989 | 1.2198 | 15.4593 | 20.3386 | 14.2394 | 21.5584 | 6.815% | 13.63% | 20.45% | 17.004 | 18.7939
Gd, ppm | 5.3284 | 0.61409 | 4.1002 | 6.5566 | 3.4861 | 7.1707 | 11.52% | 23.05% | 34.57% | 5.062 | 5.5948
Ho, ppm | 0.95534 | 0.1303 | 0.69474 | 1.2159 | 0.56444 | 1.3462 | 13.64% | 27.28% | 40.92% | 0.90757 | 1.0031
K, wt.% | 2.3788 | 0.050377 | 2.278 | 2.4796 | 2.2277 | 2.5299 | 2.118% | 4.235% | 6.353% | 2.2599 | 2.4977
La, ppm | 32.5111 | 2.668 | 27.1751 | 37.8471 | 24.5071 | 40.5151 | 8.206% | 16.41% | 24.62% | 30.8856 | 34.1367
Li, ppm | 59.0658 | 3.3923 | 52.2813 | 65.8503 | 48.889 | 69.2426 | 5.743% | 11.49% | 17.23% | 56.1125 | 62.0191
Mg, wt.% | 1.1994 | 0.03532 | 1.1287 | 1.27 | 1.0934 | 1.3053 | 2.945% | 5.89% | 8.835% | 1.1394 | 1.2593
Mn, wt.% | 0.04561 | 0.0018252 | 0.041959 | 0.04926 | 0.040134 | 0.051085 | 4.002% | 8.003% | 12.01% | 0.043329 | 0.04789
Nd, ppm | 30.4628 | 1.6511 | 27.1606 | 33.765 | 25.5095 | 35.416 | 5.42% | 10.84% | 16.26% | 28.9396 | 31.9859
Ni, ppm | 36.4081 | 6.1208 | 24.1664 | 48.6497 | 18.0456 | 54.7705 | 16.81% | 33.62% | 50.44% | 34.5877 | 38.2285
P, wt.% | 0.05748 | 0.0026857 | 0.052109 | 0.062852 | 0.049423 | 0.065537 | 4.672% | 9.345% | 14.02% | 0.054606 | 0.060354
Pr, ppm | 7.8221 | 0.77815 | 6.2658 | 9.3784 | 5.4877 | 10.1566 | 9.948% | 19.9% | 29.84% | 7.431 | 8.2132
Rb, ppm | 151.2258 | 4.5039 | 142.2181 | 160.2335 | 137.7142 | 164.7374 | 2.978% | 5.956% | 8.935% | 143.6645 | 158.7871
S, wt.% | 1.2478 | 0.021307 | 1.2052 | 1.2904 | 1.1839 | 1.3117 | 1.708% | 3.415% | 5.123% | 1.1854 | 1.3102
Si, wt.% | 25.8014 | 0.98686 | 23.8277 | 27.7751 | 22.8408 | 28.762 | 3.825% | 7.65% | 11.47% | 24.5113 | 27.0915
Sm, ppm | 6.1289 | 0.43716 | 5.2545 | 7.0032 | 4.8174 | 7.4403 | 7.133% | 14.27% | 21.4% | 5.8224 | 6.4353
Sn, wt.% | 0.23755 | 0.017856 | 0.20184 | 0.27326 | 0.18398 | 0.29112 | 7.517% | 15.03% | 22.55% | 0.22567 | 0.24943
Sr, ppm | 131.2679 | 3.0435 | 125.1809 | 137.3549 | 122.1374 | 140.3984 | 2.319% | 4.637% | 6.956% | 124.7045 | 137.8313
Ta, ppm | 0.99606 | 0.24668 | 0.5027 | 1.4894 | 0.25601 | 1.7361 | 24.77% | 49.53% | 74.3% | 0.94626 | 1.0459
Tb, ppm | 0.80155 | 0.10705 | 0.58746 | 1.0156 | 0.48041 | 1.1227 | 13.35% | 26.71% | 40.06% | 0.76147 | 0.84163
Th, ppm | 13.0093 | 0.79102 | 11.4273 | 14.5914 | 10.6363 | 15.3824 | 6.08% | 12.16% | 18.24% | 12.3589 | 13.6598
Ti, wt.% | 0.34649 | 0.011606 | 0.32327 | 0.3697 | 0.31167 | 0.3813 | 3.35% | 6.699% | 10.05% | 0.32916 | 0.36381
Tm, ppm | 0.37374 | 0.057656 | 0.25843 | 0.48905 | 0.20077 | 0.54671 | 15.43% | 30.85% | 46.28% | 0.35505 | 0.39243
U, ppm | 3.0795 | 0.31258 | 2.4543 | 3.7047 | 2.1418 | 4.0172 | 10.15% | 20.3% | 30.45% | 2.9255 | 3.2335
V, ppm | 69.9663 | 8.092 | 53.7824 | 86.1503 | 45.6905 | 94.2422 | 11.57% | 23.13% | 34.7% | 66.468 | 73.4646
Y, ppm | 25.6143 | 2.3464 | 20.9214 | 30.3071 | 18.575 | 32.6535 | 9.161% | 18.32% | 27.48% | 24.3336 | 26.895
Yb, ppm | 2.42 | 0.47645 | 1.4671 | 3.3729 | 0.99061 | 3.8493 | 19.69% | 39.38% | 59.07% | 2.299 | 2.541
Zn, ppm | 107.5281 | 15.5827 | 76.3626 | 138.6935 | 60.7799 | 154.2762 | 14.49% | 28.98% | 43.48% | 102.1517 | 112.9045
4-Acid Digestion
Ag, ppm | 0.49727 | 0.025744 | 0.44579 | 0.54876 | 0.42004 | 0.57451 | 5.177% | 10.35% | 15.53% | 0.47241 | 0.52214
Al, wt.% | 6.3685 | 0.20019 | 5.9681 | 6.7688 | 5.7679 | 6.969 | 3.143% | 6.287% | 9.43% | 6.05 | 6.6869
As, wt.% | 0.116 | 0.0034555 | 0.10908 | 0.12291 | 0.10563 | 0.12636 | 2.979% | 5.958% | 8.937% | 0.1102 | 0.1218
Ba, ppm | 655.7836 | 16.8585 | 622.0666 | 689.5005 | 605.2081 | 706.359 | 2.571% | 5.141% | 7.712% | 622.9944 | 688.5727
Be, ppm | 3.0464 | 0.18849 | 2.6694 | 3.4233 | 2.4809 | 3.6118 | 6.187% | 12.37% | 18.56% | 2.894 | 3.1987
Bi, ppm | 23.3702 | 1.433 | 20.5041 | 26.2363 | 19.071 | 27.6693 | 6.132% | 12.26% | 18.4% | 22.2016 | 24.5387
Ca, wt.% | 6.8532 | 0.1662 | 6.5208 | 7.1856 | 6.3546 | 7.3519 | 2.425% | 4.85% | 7.276% | 6.5106 | 7.1959
Ce, ppm | 67.0281 | 3.8118 | 59.4046 | 74.6517 | 55.5928 | 78.4635 | 5.687% | 11.37% | 17.06% | 63.6767 | 70.3795
Co, ppm | 13.0289 | 0.85601 | 11.3169 | 14.7409 | 10.4609 | 15.597 | 6.57% | 13.14% | 19.71% | 12.3775 | 13.6804
Cr, ppm | 67.9856 | 8.0707 | 51.8443 | 84.127 | 43.7736 | 92.1976 | 11.87% | 23.74% | 35.61% | 64.5863 | 71.3849
Cs, ppm | 9.3066 | 0.51971 | 8.2672 | 10.346 | 7.7474 | 10.8657 | 5.584% | 11.17% | 16.75% | 8.8412 | 9.7719
Cu, ppm | 327.7761 | 6.9549 | 313.8663 | 341.6859 | 306.9114 | 348.6408 | 2.122% | 4.244% | 6.366% | 311.3873 | 344.1649
Dy, ppm | 3.0797 | 0.27092 | 2.5379 | 3.6216 | 2.267 | 3.8925 | 8.797% | 17.59% | 26.39% | 2.9258 | 3.2337
Er, ppm | 1.4582 | 0.1953 | 1.0676 | 1.8488 | 0.87225 | 2.0441 | 13.39% | 26.79% | 40.18% | 1.3853 | 1.5311
Eu, ppm | 1.1124 | 0.045466 | 1.0214 | 1.2033 | 0.97596 | 1.2488 | 4.087% | 8.175% | 12.26% | 1.0567 | 1.168
Fe, wt.% | 4.7907 | 0.0926 | 4.6055 | 4.9759 | 4.5129 | 5.0685 | 1.933% | 3.866% | 5.799% | 4.5511 | 5.0302
Ga, ppm | 17.8847 | 0.84728 | 16.1901 | 19.5792 | 15.3428 | 20.4265 | 4.737% | 9.475% | 14.21% | 16.9904 | 18.7789
Gd, ppm | 4.626 | 0.4178 | 3.7904 | 5.4616 | 3.3727 | 5.8794 | 9.031% | 18.06% | 27.09% | 4.3947 | 4.8573
Hf, ppm | 2.6627 | 0.20708 | 2.2485 | 3.0768 | 2.0415 | 3.2839 | 7.777% | 15.55% | 23.33% | 2.5296 | 2.7958
Ho, ppm | 0.52131 | 0.065818 | 0.38968 | 0.65295 | 0.32386 | 0.71877 | 12.63% | 25.25% | 37.88% | 0.49525 | 0.54738
In, ppm | 0.87151 | 0.21769 | 0.43613 | 1.3069 | 0.21844 | 1.5246 | 24.98% | 49.96% | 74.94% | 0.82793 | 0.91508
K, wt.% | 2.3554 | 0.06461 | 2.2261 | 2.4846 | 2.1615 | 2.5492 | 2.743% | 5.486% | 8.229% | 2.2376 | 2.4731
La, ppm | 33.2382 | 2.0385 | 29.1612 | 37.3151 | 27.1227 | 39.3536 | 6.133% | 12.27% | 18.4% | 31.5763 | 34.9001
Li, ppm | 58.9647 | 3.07 | 52.8246 | 65.1047 | 49.7546 | 68.1747 | 5.207% | 10.41% | 15.62% | 56.0164 | 61.9129
Lu, ppm | 0.19187 | 0.017557 | 0.15675 | 0.22698 | 0.1392 | 0.24454 | 9.151% | 18.3% | 27.45% | 0.18227 | 0.20146
Mg, wt.% | 1.1509 | 0.02683 | 1.0972 | 1.2045 | 1.0704 | 1.2313 | 2.331% | 4.663% | 6.994% | 1.0933 | 1.2084
Mn, wt.% | 0.044617 | 0.0014874 | 0.041642 | 0.047591 | 0.040154 | 0.049079 | 3.334% | 6.668% | 10% | 0.042386 | 0.046847
Mo, ppm | 1.3203 | 0.092544 | 1.1352 | 1.5054 | 1.0427 | 1.598 | 7.009% | 14.02% | 21.03% | 1.2543 | 1.3863
Na, wt.% | 0.92567 | 0.034687 | 0.85629 | 0.99504 | 0.82161 | 1.0297 | 3.747% | 7.494% | 11.24% | 0.87938 | 0.97195
Nb, ppm | 11.2766 | 1.0377 | 9.2012 | 13.3519 | 8.1635 | 14.3896 | 9.202% | 18.4% | 27.61% | 10.7127 | 11.8404
Nd, ppm | 28.28 | 2.5434 | 23.1932 | 33.3668 | 20.6498 | 35.9102 | 8.994% | 17.99% | 26.98% | 26.866 | 29.694
Ni, ppm | 33.8525 | 2.3718 | 29.1089 | 38.5961 | 26.7372 | 40.9679 | 7.006% | 14.01% | 21.02% | 32.1599 | 35.5451
P, wt.% | 0.057834 | 0.00171 | 0.054414 | 0.061254 | 0.052704 | 0.062964 | 2.957% | 5.914% | 8.87% | 0.054942 | 0.060726
Pb, ppm | 27.0692 | 1.5854 | 23.8985 | 30.2399 | 22.3131 | 31.8252 | 5.857% | 11.71% | 17.57% | 25.7157 | 28.4226
Pr, ppm | 8.0388 | 0.47613 | 7.0866 | 8.9911 | 6.6104 | 9.4672 | 5.923% | 11.85% | 17.77% | 7.6369 | 8.4408
Rb, ppm | 146.0925 | 8.0065 | 130.0795 | 162.1055 | 122.0731 | 170.1119 | 5.48% | 10.96% | 16.44% | 138.7879 | 153.3971
S, wt.% | 1.2245 | 0.036395 | 1.1517 | 1.2973 | 1.1153 | 1.3337 | 2.972% | 5.944% | 8.917% | 1.1633 | 1.2857
Sb, ppm | 1.5409 | 0.1867 | 1.1675 | 1.9143 | 0.9808 | 2.101 | 12.12% | 24.23% | 36.35% | 1.4639 | 1.6179
Sc, ppm | 10.108 | 0.43703 | 9.2339 | 10.982 | 8.7969 | 11.419 | 4.324% | 8.647% | 12.97% | 9.6026 | 10.6134
Sm, ppm | 5.6545 | 0.33829 | 4.9779 | 6.331 | 4.6396 | 6.6693 | 5.983% | 11.97% | 17.95% | 5.3717 | 5.9372
Sr, ppm | 128.0104 | 4.6076 | 118.7951 | 137.2256 | 114.1875 | 141.8333 | 3.599% | 7.199% | 10.8% | 121.6098 | 134.4109
Ta, ppm | 0.95618 | 0.11231 | 0.73156 | 1.1808 | 0.61925 | 1.2931 | 11.75% | 23.49% | 35.24% | 0.90837 | 1.004
Tb, ppm | 0.60025 | 0.035942 | 0.52837 | 0.67214 | 0.49243 | 0.70808 | 5.988% | 11.98% | 17.96% | 0.57024 | 0.63026
Th, ppm | 12.9408 | 0.76798 | 11.4048 | 14.4768 | 10.6369 | 15.2448 | 5.935% | 11.87% | 17.8% | 12.2938 | 13.5878
Ti, wt.% | 0.32478 | 0.017451 | 0.28988 | 0.35968 | 0.27243 | 0.37713 | 5.373% | 10.75% | 16.12% | 0.30854 | 0.34102
Tl, ppm | 0.78852 | 0.026875 | 0.73477 | 0.84226 | 0.70789 | 0.86914 | 3.408% | 6.816% | 10.22% | 0.74909 | 0.82794
Tm, ppm | 0.17964 | 0.017834 | 0.14397 | 0.21531 | 0.12613 | 0.23314 | 9.928% | 19.86% | 29.78% | 0.17066 | 0.18862
U, ppm | 2.8044 | 0.25127 | 2.3019 | 3.3069 | 2.0506 | 3.5582 | 8.96% | 17.92% | 26.88% | 2.6642 | 2.9446
V, ppm | 71.734 | 1.7512 | 68.2317 | 75.2364 | 66.4805 | 76.9876 | 2.441% | 4.882% | 7.324% | 68.1473 | 75.3207
W, ppm | 21.3278 | 1.0919 | 19.1439 | 23.5117 | 18.0519 | 24.6036 | 5.12% | 10.24% | 15.36% | 20.2614 | 22.3941
Y, ppm | 13.7663 | 1.3751 | 11.0161 | 16.5165 | 9.641 | 17.8915 | 9.989% | 19.98% | 29.97% | 13.078 | 14.4546
Yb, ppm | 1.2672 | 0.1139 | 1.0394 | 1.495 | 0.92552 | 1.6089 | 8.988% | 17.98% | 26.96% | 1.2038 | 1.3306
Zn, ppm | 99.8651 | 4.3887 | 91.0876 | 108.6425 | 86.6989 | 113.0313 | 4.395% | 8.789% | 13.18% | 94.8718 | 104.8583
Zr, ppm | 86.0146 | 7.9397 | 70.1353 | 101.8939 | 62.1956 | 109.8336 | 9.231% | 18.46% | 27.69% | 81.7139 | 90.3153
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 143
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 143
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Intertek, Perth, WA, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Ankara, Anatolia, Turkey
SGS Australia Mineral Services, Perth, WA, Australia
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Fusion XRF
| Analytical results for Al2O3 in OREAS 143 (Certified Value 12.38 wt.%) | [15 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 12.2 | 11.52 | 12.471 | 12.12 | 11.95 | 12.15 | 12.4 | 12.45 | 12.2 | 12.8182 | 12.4937 | 12.22 | 12.82 | 12.5436 | 1
| 1 | 2 | 12.19 | 11.81 | 12.376 | 12.14 | 12.4 | 12.3 | 12.4 | 12.5 | 12.35 | 12.738 | 12.5057 | 12.28 | 13.48 | 12.4096 | #N/A
| 1 | 3 | 12.23 | 11.74 | 12.452 | 12.12 | 11.84 | 12.15 | 12.45 | 12.4 | 12.3 | 12.7674 | 12.3237 | 12.3 | 11.81 | 12.3187 | 16
| 1 | 4 | 12.18 | 11.8 | 12.527 | 12.11 | 11.77 | 12.15 | 12.45 | 12.6 | 12.25 | 12.6791 | 12.2193 | 12.3 | 12.58 | 12.5846 | 12.3832
| 1 | 5 | 12.21 | 11.77 | 12.395 | 12.08 | 11.65 | 12.25 | 12.45 | 12.45 | 12.25 | 12.821 | 12.4325 | 12.34 | 13.3 | 12.4438 | 7
| 1 | 6 | 12.19 | 11.99 | 12.527 | 12.12 | 11.19 | 12.25 | 12.4 | 12.55 | 12.15 | 12.8197 | 12.3444 | 12.38 | 12.94 | 12.393
| Mean |  | 12.2 | 11.7717 | 12.458 | 12.115 | 11.8 | 12.2083 | 12.425 | 12.4917 | 12.25 | 12.7739 | 12.3866 | 12.3033 | 12.8217 | 12.4489
| Median |  | 12.195 | 11.785 | 12.4615 | 12.12 | 11.805 | 12.2 | 12.425 | 12.475 | 12.25 | 12.7928 | 12.3885 | 12.3 | 12.88 | 12.4267
| Std Dev. |  | 0.017889 | 0.15118 | 0.063906 | 0.019748 | 0.39487 | 0.066458 | 0.027386 | 0.073598 | 0.070711 | 0.057603 | 0.11089 | 0.054283 | 0.59297 | 0.099047
| Rel.Std.Dev. |  | 0.1466% | 1.284% | 0.513% | 0.163% | 3.346% | 0.5444% | 0.2204% | 0.5892% | 0.5772% | 0.4509% | 0.8952% | 0.4412% | 4.625% | 0.7956%
| PDM3 |  | -1.479% | -4.938% | 0.6041% | -2.166% | -4.71% | -1.412% | 0.3376% | 0.876% | -1.076% | 3.155% | 0.02725% | -0.6449% | 3.541% | 0.5305%
| Z-Score (Absolute) |  | 0.69 | 2.75 | 0.54 | 1.1 | 2.61 | 0.65 | 0.38 | 0.7 | 0.45 | 2.05 | 0.2 | 0.2 | 2.28 | 0.5
| Analytical results for As in OREAS 143 (Certified Value 0.114 wt.%) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 15 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.13 | 0.11 | 0.106 | 0.101 | 0.14 | 0.1 | 0.11 | 0.12 | 0.12 | [5 empty] | 1
| 1 | 2 | 0.133 | 0.1 | 0.111 | 0.103 | 0.14 | 0.11 | 0.11 | 0.12 | 0.12 | [5 empty] | #N/A
| 1 | 3 | 0.129 | 0.11 | 0.108 | 0.102 | 0.14 | 0.12 | 0.11 | 0.12 | 0.12 | [5 empty] | 16
| 1 | 4 | 0.129 | 0.12 | 0.109 | 0.105 | 0.13 | 0.1 | 0.11 | 0.12 | 0.12 | [5 empty] | 0.11372
| 1 | 5 | 0.129 | 0.11 | 0.11 | 0.106 | 0.14 | 0.1 | 0.11 | 0.12 | 0.12 | [5 empty] | 8
| 1 | 6 | 0.13 | 0.11 | 0.106 | 0.105 | 0.14 | 0.11 | 0.11 | 0.12 | 0.13
| Mean |  | 0.13 | 0.11 | 0.10833 | 0.10367 | 0.13833 | 0.10667 | 0.11 | 0.12 | 0.12167
| Median |  | 0.1295 | 0.11 | 0.1085 | 0.104 | 0.14 | 0.105 | 0.11 | 0.12 | 0.12
| Std Dev. |  | 0.0015492 | 0.0063246 | 0.0020656 | 0.0019664 | 0.0040825 | 0.008165 | 0 | 0 | 0.0040825
| Rel.Std.Dev. |  | 1.192% | 5.75% | 1.907% | 1.897% | 2.951% | 7.655% | 0% | 0% | 3.355%
| PDM3 |  | 14.32% | -3.268% | -4.734% | -8.838% | 21.65% | -6.2% | -3.268% | 5.525% | 6.991%
| Z-Score (Absolute) |  | 2.13 | 0 | 0.18 | 0.67 | 3.02 | 0.35 | 0 | 1.06 | 1.24
| Analytical results for BaO in OREAS 143 (Certified Value 738 ppm) | [15 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | 983 | 770 | 600 | 700 | 700 | 800 | 728.293 | 764.9575 | 800 | 500 | 850 |  |  | 1
| 1 | 2 | < 100 | 927 | 750 | 700 | 700 | 800 | 800 | 730.9391 | 784.4722 | 800 | 600 | 860 |  |  | #N/A
| 1 | 3 | < 100 | 983 | 750 | 600 | 700 | 700 | 700 | 721.5493 | 776.8319 | 700 | 300 | 810 |  |  | 16
| 1 | 4 | < 100 | 1016 | 730 | 600 | 700 | 700 | 700 | 725.7697 | 784.7555 | 800 | 600 | 1010 |  |  | 737.7848
| 1 | 5 | < 100 | 927 | 720 | 600 | 700 | 800 | 700 | 719.6066 | 753.9116 | 900 | 600 | 810 |  |  | 9
| 1 | 6 | < 100 | 960 | 740 | 600 | 700 | 700 | 700 | 723.76 | 773.5303 | 700 | 700 | 830
| Mean |  |  | 966 | 743.3333 | 616.6667 | 700 | 733.3333 | 733.3333 | 724.9862 | 773.0765 | 783.3333 | 550 | 861.6667
| Median |  |  | 971.5 | 745 | 600 | 700 | 700 | 700 | 724.7648 | 775.1811 | 800 | 600 | 840
| Std Dev. |  |  | 35.0885 | 17.5119 | 40.8248 | 0 | 51.6398 | 51.6398 | 4.2247 | 11.9481 | 75.2773 | 137.8405 | 75.4763
| Rel.Std.Dev. |  |  | 3.632% | 2.356% | 6.62% | 0% | 7.042% | 7.042% | 0.5827% | 1.546% | 9.61% | 25.06% | 8.759%
| PDM3 |  |  | 30.93% | 0.7521% | -16.42% | -5.121% | -0.6033% | -0.6033% | -1.735% | 4.783% | 6.174% | -25.45% | 16.79%
| Z-Score (Absolute) |  | 10.27 | 3.49 | 0.15 | 1.75 | 0.5 | 0 | 0 | 0.13 | 0.6 | 0.75 | 2.76 | 1.93
Results from laboratory 17 were removed due to their high Z-score and unsatisfactory performance on the control sample.
| Analytical results for Bi in OREAS 143 (Indicative Value 81 ppm) | [15 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 09 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | 59 | 100 | < 100 | < 100 | < 100 | [8 empty] | 1
| 1 | 2 | < 100 | 64 | 100 | < 100 | < 100 | < 100 | [8 empty] | 1
| 1 | 3 | < 100 | 58 | 100 | < 100 | < 100 | < 100 | [8 empty] | 16
| 1 | 4 | < 100 | 64 | 100 | < 100 | < 100 | < 100 | [8 empty] | 80.9167
| 1 | 5 | < 100 | 58 | 100 | < 100 | < 100 | < 100 | [8 empty] | 7
| 1 | 6 | < 100 | 68 | 100 | < 100 | < 100 | < 100
| Mean |  |  | 61.8333 | 100
| Median |  |  | 61.5 | 100
| Std Dev. |  |  | 4.1191 | 0
| Rel.Std.Dev. |  |  | 6.662% | 0%
| PDM3 |  |  | -23.58% | 23.58%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 143 (Certified Value 9.79 wt.%) | [15 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 9.87 | 12.34 | 9.8686 | 9.69 | 9.71 | 9.59 | 9.52 | 9.74 | 9.74 | 10.1576 | 10.0369 | 9.54 | 10.05 | 10.021 | 1
| 1 | 2 | 9.68 | 12.3 | 9.8798 | 9.69 | 9.58 | 9.63 | 9.65 | 9.79 | 9.75 | 10.1298 | 9.9055 | 9.74 | 9.97 | 9.729 | #N/A
| 1 | 3 | 9.78 | 12.44 | 9.8588 | 9.77 | 9.68 | 9.74 | 9.7 | 9.82 | 9.74 | 10.1363 | 9.808 | 9.6 | 8.79 | 9.689 | 16
| 1 | 4 | 9.73 | 12.35 | 9.9063 | 9.7 | 9.9 | 9.7 | 9.57 | 9.74 | 9.86 | 10.1635 | 10.0893 | 9.58 | 9.62 | 9.902 | 9.7931
| 1 | 5 | 9.73 | 12.4 | 9.814 | 9.62 | 9.82 | 9.71 | 9.5 | 9.72 | 9.77 | 10.0938 | 9.9055 | 9.68 | 10.07 | 9.884 | 10
| 1 | 6 | 9.69 | 12.32 | 9.8937 | 9.8 | 9.73 | 9.65 | 9.79 | 9.92 | 9.61 | 10.085 | 9.893 | 9.57 | 9.85 | 9.708
| Mean |  | 9.7467 | 12.3583 | 9.8702 | 9.7117 | 9.7367 | 9.67 | 9.6217 | 9.7883 | 9.745 | 10.1277 | 9.9397 | 9.6183 | 9.725 | 9.8222
| Median |  | 9.73 | 12.345 | 9.8742 | 9.695 | 9.72 | 9.675 | 9.61 | 9.765 | 9.745 | 10.1331 | 9.9055 | 9.59 | 9.91 | 9.8065
| Std Dev. |  | 0.070048 | 0.052313 | 0.032368 | 0.064317 | 0.11147 | 0.056214 | 0.11232 | 0.074409 | 0.080187 | 0.032333 | 0.10364 | 0.076004 | 0.48694 | 0.13355
| Rel.Std.Dev. |  | 0.7187% | 0.4233% | 0.3279% | 0.6623% | 1.145% | 0.5813% | 1.167% | 0.7602% | 0.8229% | 0.3193% | 1.043% | 0.7902% | 5.007% | 1.36%
| PDM3 |  | -0.474% | 26.19% | 0.7875% | -0.8314% | -0.5761% | -1.257% | -1.75% | -0.04849% | -0.491% | 3.417% | 1.497% | -1.784% | -0.6952% | 0.297%
| Z-Score (Absolute) |  | 0.01 | 23.15 | 1.1 | 0.3 | 0.08 | 0.67 | 1.1 | 0.38 | 0.01 | 3.38 | 1.72 | 1.13 | 0.18 | 0.68
| Analytical results for Cd in OREAS 143 (Indicative Value < 10 ppm) | [15 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 02 | 03 | [12 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [12 empty] | 1
| 1 | 2 | < 100 | < 10 | [12 empty] | 2
| 1 | 3 | < 100 | < 10 | [12 empty] | 16
| 1 | 4 | < 100 | < 10 | [12 empty] | < 10
| 1 | 5 | < 100 | < 10 | [12 empty] | 8
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 143 (Indicative Value < 80 ppm) | [15 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 06 | 07 | 08 | 09 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 80 | 80 | < 80 | < 80 | [10 empty] | 1
| 1 | 2 | < 80 | < 80 | 80 | < 80 | [10 empty] | 3
| 1 | 3 | < 80 | < 80 | < 80 | < 80 | [10 empty] | 16
| 1 | 4 | < 80 | < 80 | 80 | < 80 | [10 empty] | < 80
| 1 | 5 | < 80 | < 80 | < 80 | 80 | [10 empty] | 9
| 1 | 6 | < 80 | 80 | < 80 | < 80
| Mean |  |  | 80 | 80 | 80
| Median |  |  | 80 | 80 | 80
| Std Dev. |  |  | 0 | 0
| Rel.Std.Dev. |  |  | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 2.02 | 0.67 | 0.67 | 0.67
| Analytical results for Cl in OREAS 143 (Indicative Value < 50 ppm) | [15 empty] | Indicative
Cl | Round | Replicate | Lab | [13 empty] | 1
| No | No | 04 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 50 | [13 empty] | 1
| 1 | 2 | < 50 | [13 empty] | 4
| 1 | 3 | < 50 | [13 empty] | 16
| 1 | 4 | < 50 | [13 empty] | < 50
| 1 | 5 | < 50 | [13 empty] | 10
| 1 | 6 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 143 (Certified Value < 100 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | [6 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 37 | < 100 | < 100 | < 100 | < 100 | [6 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | 59 | < 100 | < 100 | < 100 | < 100 | [6 empty] | #N/A
| 1 | 3 | < 100 | < 50 | < 50 | 51 | < 100 | < 100 | < 100 | < 100 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | 59 | < 100 | < 100 | < 100 | < 100 | [6 empty] | < 100
| 1 | 5 | < 100 | < 50 | < 50 | 37 | < 100 | < 100 | < 100 | < 100 | [6 empty] | 11
| 1 | 6 | < 100 | < 50 | < 50 | 66 | < 100 | < 100 | < 100 | < 100
| Mean |  |  |  |  | 54.4
| Median |  |  |  |  | 59
| Std Dev. |  |  |  |  | 11.0815
| Rel.Std.Dev. |  |  |  |  | 20.37%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cr2O3 in OREAS 143 (Indicative Value 137 ppm) | [15 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 14 | 15 | 17 | 18 |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | 132 | < 50 | 100 | < 150 | 150 | < 150 | 150 | 101.0586 | 100 | 90 | 120 |  |  | 1
| 1 | 2 | < 100 | 132 | < 50 | < 100 | < 150 | 150 | < 150 | 150 | 117.769 | 100 | 120 | 110 |  |  | 5
| 1 | 3 | < 100 | 161 | < 50 | < 100 | < 150 | 150 | < 150 | 150 | 118.1565 | 100 | < 50 | 100 |  |  | 16
| 1 | 4 | < 100 | 146 | < 50 | < 100 | < 150 | 150 | < 150 | 150 | 143.0568 | 100 | 50 | 130 |  |  | 137.4072
| 1 | 5 | < 100 | 161 | < 50 | 200 | < 150 | 150 | 150 | 150 | 126.1816 | 100 | < 50 | 120 |  |  | 11
| 1 | 6 | < 100 | 161 | < 50 | < 100 | < 150 | 150 | < 150 | 150 | 117.6543 | 100 | 110 | 120
| Mean |  |  | 148.8333 |  | 150 |  | 150 | 150 | 150 | 120.6461 | 100 | 92.5 | 116.6667
| Median |  |  | 153.5 |  | 150 |  | 150 | 150 | 150 | 117.9627 | 100 | 100 | 120
| Std Dev. |  |  | 14.2747 |  | 70.7107 |  | 0 |  | 0 | 13.7112 | 0 | 30.957 | 10.328
| Rel.Std.Dev. |  |  | 9.591% |  | 47.14% |  | 0% |  | 0% | 11.36% | 0% | 33.47% | 8.853%
| PDM3 |  |  | 8.316% |  | 9.165% |  | 9.165% | 9.165% | 9.165% | -12.2% | -27.22% | -32.68% | -15.09%
| Z-Score (Absolute) |  | 1.1 | 1.45 | 1.75 | 0.24 | 0.51 | 1.39 | 0.19 | 1.39 | 0.73 | 0.19 | 0.59 | 0.62
| Analytical results for Cs in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Cs | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 6
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 12
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 143 (Certified Value 319 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 14 | 15 | 17 | 18 |  | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 310 | 300 | 340 | 320 | 304 | 320 | 300 | 290 | 280 | 308.5704 | 344 | 320 | 336 |  | 1
| 1 | 2 | 340 | 300 | 340 | 300 | 304 | 330 | 320 | 260 | 300 | 329.727 | 336 | 360 | 312 |  | #N/A
| 1 | 3 | 290 | 300 | 340 | 300 | 336 | 330 | 380 | 300 | 300 | 323.6203 | 320 | 340 | 312 |  | 16
| 1 | 4 | 310 | 300 | 340 | 320 | 296 | 310 | 300 | 300 | 310 | 337.0166 | 272 | 320 | 344 |  | 319.2899
| 1 | 5 | 290 | 300 | 340 | 330 | 280 | 320 | 320 | 300 | 340 | 394.2404 | 320 | 350 | 328 |  | 12
| 1 | 6 | 300 | 300 | 340 | 320 | 256 | 300 | 320 | 290 | 300 | 393.8549 | 344 | 390 | 328
| Mean |  | 306.6667 | 300 | 340 | 315 | 296 | 318.3333 | 323.3333 | 290 | 305 | 347.8383 | 322.6667 | 346.6667 | 326.6667
| Median |  | 305 | 300 | 340 | 320 | 300 | 320 | 320 | 295 | 300 | 333.3718 | 328 | 345 | 328
| Std Dev. |  | 18.619 | 0 | 0 | 12.2474 | 26.7731 | 11.6905 | 29.4392 | 15.4919 | 19.7484 | 36.9981 | 27.09 | 26.5832 | 12.8167
| Rel.Std.Dev. |  | 6.071% | 0% | 0% | 3.888% | 9.045% | 3.672% | 9.105% | 5.342% | 6.475% | 10.64% | 8.396% | 7.668% | 3.923%
| PDM3 |  | -3.954% | -6.042% | 6.486% | -1.344% | -7.294% | -0.2996% | 1.266% | -9.173% | -4.476% | 8.941% | 1.058% | 8.574% | 2.31%
| Z-Score (Absolute) |  | 0.64 | NA | 0.92 | 0.25 | 1.15 | 0.09 | 0.14 | 1.41 | 0.71 | 1.29 | 0.09 | 1.23 | 0.28
Results from laboratory 2 were removed due to their 100 ppm reading resolution.
| Analytical results for Fe2O3 in OREAS 143 (Certified Value 6.98 wt.%) | [15 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 6.89 | 7.28 | 6.72 | 7.277 | 6.95 | 6.805 | 6.805 | 6.863 | 6.82 | 6.9473 | 7.0938 | 7.05 | 7.11 | 7.137 | 1
| 1 | 2 | 6.85 | 7.22 | 6.705 | 7.349 | 6.98 | 6.863 | 6.82 | 6.891 | 6.834 | 6.9306 | 7.1217 | 7.02 | 7.49 | 6.95 | #N/A
| 1 | 3 | 6.92 | 7.32 | 6.72 | 7.377 | 6.96 | 6.791 | 6.905 | 6.834 | 6.805 | 6.9535 | 7.0002 | 7.04 | 6.61 | 6.926 | 16
| 1 | 4 | 6.95 | 7.21 | 6.805 | 7.334 | 7.08 | 6.82 | 6.834 | 6.92 | 6.777 | 6.9718 | 7.0799 | 6.98 | 7.02 | 7.122 | 6.9844
| 1 | 5 | 6.87 | 7.28 | 6.677 | 7.363 | 6.94 | 6.891 | 6.82 | 6.848 | 6.82 | 6.9103 | 7.1123 | 7.03 | 7.37 | 7.001 | 13
| 1 | 6 | 6.89 | 7.3 | 6.834 | 7.377 | 7.03 | 6.805 | 6.891 | 6.877 | 6.748 | 6.9322 | 7.012 | 6.99 | 7.24 | 6.993
| Mean |  | 6.895 | 7.2683 | 6.7435 | 7.3462 | 6.99 | 6.8292 | 6.8458 | 6.8722 | 6.8007 | 6.941 | 7.07 | 7.0183 | 7.14 | 7.0215
| Median |  | 6.89 | 7.28 | 6.72 | 7.356 | 6.97 | 6.8125 | 6.827 | 6.87 | 6.8125 | 6.9398 | 7.0868 | 7.025 | 7.175 | 6.997
| Std Dev. |  | 0.035637 | 0.044008 | 0.061614 | 0.037749 | 0.054406 | 0.039184 | 0.041672 | 0.030954 | 0.032284 | 0.02133 | 0.051697 | 0.027869 | 0.31035 | 0.088211
| Rel.Std.Dev. |  | 0.5169% | 0.6055% | 0.9137% | 0.5139% | 0.7783% | 0.5738% | 0.6087% | 0.4504% | 0.4747% | 0.3073% | 0.7312% | 0.3971% | 4.347% | 1.256%
| PDM3 |  | -1.28% | 4.065% | -3.449% | 5.18% | 0.08021% | -2.223% | -1.984% | -1.607% | -2.631% | -0.622% | 1.225% | 0.4859% | 2.228% | 0.5312%
| Z-Score (Absolute) |  | 0.42 | 1.82 | 1.34 | 2.29 | 0.15 | 0.82 | 0.72 | 0.56 | 0.99 | 0.15 | 0.63 | 0.32 | 1.05 | 0.34
| Analytical results for Ga in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Ga | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 7
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 8
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 143 (Certified Value < 100 ppm) | [15 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 06 | 07 | 08 | 09 | 14 | 18 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 200 | [9 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 200 | [9 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 200 | 30 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 200 | [9 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 200 | [9 empty] | 14
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 200 | 30
| Mean | [6 empty] | 30
| Median | [6 empty] | 30
| Std Dev. | [6 empty] | 0
| Rel.Std.Dev. | [6 empty] | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Hg in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Hg | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 18
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
In | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 2
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 143 (Certified Value 2.84 wt.%) | [15 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.8 | 3.52 | 2.7067 | 2.81 | 2.84 | 2.85 | 2.84 | 2.81 | 2.85 | 2.9624 | 2.918 | 2.78 | 2.853 | 2.905 | 1
| 1 | 2 | 2.78 | 3.53 | 2.714 | 2.82 | 2.83 | 2.78 | 2.83 | 2.83 | 2.84 | 3.0089 | 2.9244 | 2.78 | 2.973 | 2.857 | #N/A
| 1 | 3 | 2.82 | 3.56 | 2.7369 | 2.82 | 2.82 | 2.86 | 2.84 | 2.82 | 2.84 | 2.999 | 2.8622 | 2.8 | 2.59 | 2.83 | 16
| 1 | 4 | 2.81 | 3.58 | 2.7308 | 2.82 | 2.83 | 2.8 | 2.84 | 2.87 | 2.84 | 2.9589 | 2.9313 | 2.8 | 2.819 | 2.921 | 2.8376
| 1 | 5 | 2.8 | 3.52 | 2.7188 | 2.81 | 2.77 | 2.86 | 2.84 | 2.83 | 2.84 | 2.98 | 2.8954 | 2.81 | 2.938 | 2.865 | 15
| 1 | 6 | 2.81 | 3.52 | 2.7079 | 2.81 | 2.8 | 2.84 | 2.85 | 2.82 | 2.83 | 3.0092 | 2.881 | 2.79 | 2.881 | 2.847
| Mean |  | 2.8033 | 3.5383 | 2.7192 | 2.815 | 2.815 | 2.8317 | 2.84 | 2.83 | 2.84 | 2.9864 | 2.9021 | 2.7933 | 2.8423 | 2.8708
| Median |  | 2.805 | 3.525 | 2.7164 | 2.815 | 2.825 | 2.845 | 2.84 | 2.825 | 2.84 | 2.9895 | 2.9067 | 2.795 | 2.867 | 2.861
| Std Dev. |  | 0.013663 | 0.025626 | 0.012321 | 0.0054772 | 0.025884 | 0.033714 | 0.0063246 | 0.020976 | 0.0063246 | 0.022618 | 0.027147 | 0.012111 | 0.13568 | 0.035057
| Rel.Std.Dev. |  | 0.4874% | 0.7242% | 0.4531% | 0.1946% | 0.9195% | 1.191% | 0.2227% | 0.7412% | 0.2227% | 0.7574% | 0.9354% | 0.4336% | 4.774% | 1.221%
| PDM3 |  | -1.209% | 24.69% | -4.174% | -0.7976% | -0.7976% | -0.2103% | 0.0834% | -0.269% | 0.0834% | 5.242% | 2.27% | -1.561% | 0.1656% | 1.17%
| Z-Score (Absolute) |  | 0.82 | 17.76 | 2.95 | 0.53 | 0.53 | 0.11 | 0.11 | 0.15 | 0.11 | 3.81 | 1.67 | 1.07 | 0.16 | 0.89
Results from laboratory 3 were removed due to their high Z-score and unsatisfactory performance on the control sample.
| Analytical results for La in OREAS 143 (Indicative Value < 90 ppm) | [15 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 06 | 07 | 08 | 09 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 90 | 90 | < 90 | < 90 | [10 empty] | 1
| 1 | 2 | < 90 | < 90 | < 90 | < 90 | [10 empty] | 3
| 1 | 3 | < 90 | < 90 | < 90 | < 90 | [10 empty] | 16
| 1 | 4 | < 90 | < 90 | < 90 | < 90 | [10 empty] | < 90
| 1 | 5 | < 90 | 90 | < 90 | < 90 | [10 empty] | 9
| 1 | 6 | < 90 | < 90 | < 90 | < 90
| Mean |  |  | 90
| Median |  |  | 90
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for MgO in OREAS 143 (Certified Value 2.05 wt.%) | [15 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 2.04 | 1.84 | 2.023 | 2.1 | 2.03 | 2.03 | 2.1 | 2.06 | 1.9794 | 2.0603 | 2.01 | 2.18 | 2.084 |  | 1
| 1 | 2 | 2.05 | 1.88 | 1.99 | 2.12 | 2.05 | 2.04 | 2.11 | 2.09 | 1.8825 | 2.0649 | 2 | 2.24 | 2.058 |  | #N/A
| 1 | 3 | 2.06 | 1.79 | 1.973 | 2.09 | 2.03 | 2.02 | 2.1 | 2.06 | 1.8653 | 2.0291 | 2.04 | 2.01 | 2.025 |  | 16
| 1 | 4 | 2.04 | 1.76 | 1.973 | 2.1 | 2.02 | 2.05 | 2.13 | 2.07 | 1.9795 | 2.039 | 2.01 | 2.18 | 2.085 |  | 2.0511
| 1 | 5 | 2.03 | 1.85 | 1.94 | 2.11 | 2.05 | 2.06 | 2.11 | 2.06 | 1.8242 | 2.052 | 2.03 | 2.23 | 2.045 |  | 16
| 1 | 6 | 2.05 | 1.77 | 1.99 | 2.11 | 2.05 | 2.04 | 2.12 | 2.04 | 1.834 | 2.053 | 2.03 | 2.21 | 2.04
| Mean |  | 2.045 | 1.815 | 1.9815 | 2.105 | 2.0383 | 2.04 | 2.1117 | 2.0633 | 1.8942 | 2.0497 | 2.02 | 2.175 | 2.0562
| Median |  | 2.045 | 1.815 | 1.9815 | 2.105 | 2.04 | 2.04 | 2.11 | 2.06 | 1.8739 | 2.0525 | 2.02 | 2.195 | 2.0515
| Std Dev. |  | 0.010488 | 0.048477 | 0.027326 | 0.010488 | 0.013292 | 0.014142 | 0.01169 | 0.01633 | 0.069313 | 0.013394 | 0.015492 | 0.084558 | 0.024359
| Rel.Std.Dev. |  | 0.5129% | 2.671% | 1.379% | 0.4982% | 0.6521% | 0.6932% | 0.5536% | 0.7914% | 3.659% | 0.6535% | 0.7669% | 3.888% | 1.185%
| PDM3 |  | -0.2962% | -11.51% | -3.392% | 2.629% | -0.6212% | -0.54% | 2.954% | 0.5976% | -7.651% | -0.06676% | -1.515% | 6.042% | 0.2482%
| Z-Score (Absolute) |  | 0 | 6.2 | 1.71 | 1.62 | 0.18 | 0.13 | 1.8 | 0.49 | 4.07 | 0.13 | 0.67 | 3.51 | 0.3
| Analytical results for MnO in OREAS 143 (Certified Value 0.058 wt.%) | [15 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.06 | 0.06 | 0.0529 | 0.07 | 0.074 | 0.052 | 0.052 | 0.065 | 0.065 | 0.0584 | 0.054225 | 0.047 | 0.051 | 0.063 | 1
| 1 | 2 | 0.06 | 0.05 | 0.0504 | 0.06 | 0.074 | 0.052 | 0.052 | 0.065 | 0.052 | 0.058052 | 0.05594 | 0.047 | 0.057 | 0.057 | #N/A
| 1 | 3 | 0.06 | 0.09 | 0.0516 | 0.07 | 0.072 | 0.052 | 0.065 | 0.065 | 0.065 | 0.05834 | 0.055233 | 0.047 | 0.041 | 0.058 | 16
| 1 | 4 | 0.06 | 0.08 | 0.0516 | 0.07 | 0.074 | 0.052 | 0.052 | 0.065 | 0.052 | 0.058523 | 0.056849 | 0.047 | 0.048 | 0.063 | 0.058288
| 1 | 5 | 0.06 | 0.07 | 0.0478 | 0.07 | 0.067 | 0.052 | 0.052 | 0.065 | 0.052 | 0.058384 | 0.055053 | 0.056 | 0.049 | 0.063 | 17
| 1 | 6 | 0.06 | 0.07 | 0.0516 | 0.07 | 0.065 | 0.052 | 0.052 | 0.065 | 0.065 | 0.05816 | 0.054999 | 0.047 | 0.049 | 0.06
| Mean |  | 0.06 | 0.07 | 0.050983 | 0.068333 | 0.071 | 0.052 | 0.054167 | 0.065 | 0.0585 | 0.05831 | 0.055383 | 0.0485 | 0.049167 | 0.060667
| Median |  | 0.06 | 0.07 | 0.0516 | 0.07 | 0.073 | 0.052 | 0.052 | 0.065 | 0.0585 | 0.058362 | 0.055143 | 0.047 | 0.049 | 0.0615
| Std Dev. |  | 0 | 0.014142 | 0.0017486 | 0.0040825 | 0.004 | 0 | 0.0053072 | 0 | 0.0071204 | 0.00017272 | 0.00090312 | 0.0036742 | 0.0051543 | 0.0027325
| Rel.Std.Dev. |  | 0% | 20.2% | 3.43% | 5.974% | 5.634% | 0% | 9.798% | 0% | 12.17% | 0.2962% | 1.631% | 7.576% | 10.48% | 4.504%
| PDM3 |  | 2.937% | 20.09% | -12.53% | 17.23% | 21.81% | -10.79% | -7.071% | 11.52% | 0.3635% | 0.03682% | -4.984% | -16.79% | -15.65% | 4.081%
| Z-Score (Absolute) |  | 0.19 | 1.23 | 0.75 | 1.06 | 1.34 | 0.68 | 0.46 | 0.66 | 0.01 | 0.01 | 0.29 | 1.06 | 0.94 | 0.26
| Analytical results for Mo in OREAS 143 (Certified Value < 50 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 09 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 10 | < 50 | < 50 | < 50 | < 50 | [8 empty] | 1
| 1 | 2 | < 100 | 14 | < 50 | < 50 | < 50 | < 50 | [8 empty] | 14
| 1 | 3 | < 100 | 12 | 50 | 110 | < 50 | < 50 | [8 empty] | 16
| 1 | 4 | < 100 | 19 | < 50 | < 50 | < 50 | < 50 | [8 empty] | < 50
| 1 | 5 | < 100 | 16 | 50 | < 50 | < 50 | < 50 | [8 empty] | 18
| 1 | 6 | < 100 | 15 | 70 | < 50 | < 50 | < 50
| Mean |  |  | 15.2 | 56.6667 | 110
| Median |  |  | 15 | 50 | 110
| Std Dev. |  |  | 2.5884 | 11.547
| Rel.Std.Dev. |  |  | 17.03% | 20.38%
PDM3
| Z-Score (Absolute) |  | 1.53 | 1.58 | 0.75 | 0.6 | 0.6 | 0.6
| Analytical results for Na2O in OREAS 143 (Certified Value 1.21 wt.%) | [15 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 12 | 14 | 15 | 17 | 18 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [6 empty] | 2
3
| 1 | 1 | 1.23 | 1.2 | 1.19 | 1.0018 | 1.236 | 1.71 | 1.01 | 1.327 | [6 empty] | 1
| 1 | 2 | 1.21 | 1.23 | 1.19 | 1.0025 | 1.2434 | 1.73 | 1.12 | 1.321 | [6 empty] | #N/A
| 1 | 3 | 1.26 | 1.2 | 1.19 | 1.0108 | 1.2259 | 1.73 | 0.97 | 1.316 | [6 empty] | 16
| 1 | 4 | 1.25 | 1.26 | 1.2 | 1.0024 | 1.2401 | 1.72 | 1.19 | 1.325 | [6 empty] | 1.2131
| 1 | 5 | 1.23 | 1.28 | 1.19 | 1.0147 | 1.2356 | 1.71 | 1.09 | 1.302 | [6 empty] | 19
| 1 | 6 | 1.26 | 1.23 | 1.2 | 1.0042 | 1.2344 | 1.72 | 0.98 | 1.304
| Mean |  | 1.24 | 1.2333 | 1.1933 | 1.0061 | 1.2359 | 1.72 | 1.06 | 1.3158
| Median |  | 1.24 | 1.23 | 1.19 | 1.0034 | 1.2358 | 1.72 | 1.05 | 1.3185
| Std Dev. |  | 0.02 | 0.032042 | 0.005164 | 0.0053971 | 0.0059506 | 0.0089443 | 0.087636 | 0.010647
| Rel.Std.Dev. |  | 1.613% | 2.598% | 0.4327% | 0.5364% | 0.4815% | 0.52% | 8.268% | 0.8092%
| PDM3 |  | 2.22% | 1.671% | -1.627% | -17.06% | 1.882% | 41.79% | -12.62% | 8.472%
| Z-Score (Absolute) |  | 0.06 | 0.01 | 0.45 | 2.52 | 0.01 | 5.34 | 1.92 | 0.89
| Analytical results for Nb in OREAS 143 (Indicative Value < 50 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 06 | 07 | 08 | 09 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | < 50 | [10 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 50 | [10 empty] | 4
| 1 | 3 | < 50 | < 50 | < 50 | < 50 | [10 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 50 | [10 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | < 50 | [10 empty] | 10
| 1 | 6 | < 50 | < 50 | < 50 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for NiO in OREAS 143 (Indicative Value < 64 ppm) | [15 empty] | Indicative
NiO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 14 | 15 | 18 |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 250 | < 13 | < 64 | 90 | < 64 | < 64 | < 64 | < 64 | < 250 | < 100 |  |  |  |  | 1
| 1 | 2 | 250 | < 13 | < 64 | 90 | < 64 | < 64 | < 64 | 76 | < 250 | < 100 |  |  |  |  | 5
| 1 | 3 | 510 | < 13 | < 64 | 70 | < 64 | < 64 | < 64 | 76 | < 250 | < 100 | 10 |  |  |  | 16
| 1 | 4 | 380 | < 13 | < 64 | 30 | < 64 | < 64 | < 64 | < 64 | < 250 | < 100 | 20 |  |  |  | < 64
| 1 | 5 | 380 | < 13 | < 64 | 70 | < 64 | < 64 | 102 | < 64 | < 250 | < 100 | 30 |  |  |  | 11
| 1 | 6 | 250 | < 13 | < 64 | < 30 | < 64 | < 64 | 64 | < 64 | < 250 | < 100 | 30
| Mean |  | 336.6667 |  |  | 70 |  |  | 83 | 76 |  |  | 22.5
| Median |  | 315 |  |  | 70 |  |  | 83 | 76 |  |  | 25
| Std Dev. |  | 106.1446 |  |  | 24.4949 |  |  | 26.8701 | 0 |  |  | 9.5743
| Rel.Std.Dev. |  | 31.53% |  |  | 34.99% |  |  | 32.37% | 0% |  |  | 42.55%
PDM3
| Z-Score (Absolute) |  | 13.29 | 1.83 | 0.67 | 0.64 | 0.67 | 0.67 | 0.1 | 0 | 3.56 | 0.15 | 1.1
| Analytical results for P2O5 in OREAS 143 (Certified Value 0.127 wt.%) | [15 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.129 | 0.28 | 0.0321 | 0.128 | 0.13 | 0.12 | 0.13 | 0.13 | 0.12 | 0.12567 | 0.12475 | 0.14 | 0.121 | 0.137 | 1
| 1 | 2 | 0.128 | 0.31 | 0.0298 | 0.128 | 0.11 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12643 | 0.12555 | 0.14 | 0.126 | 0.13 | #N/A
| 1 | 3 | 0.127 | 0.24 | 0.0367 | 0.127 | 0.12 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12586 | 0.12597 | 0.14 | 0.121 | 0.129 | 16
| 1 | 4 | 0.129 | 0.24 | 0.0367 | 0.127 | 0.14 | 0.12 | 0.13 | 0.13 | 0.13 | 0.12583 | 0.12843 | 0.13 | 0.12 | 0.14 | 0.12658
| 1 | 5 | 0.13 | 0.25 | 0.0367 | 0.127 | 0.13 | 0.12 | 0.13 | 0.13 | 0.12 | 0.12632 | 0.12933 | 0.14 | 0.14 | 0.135 | 20
| 1 | 6 | 0.13 | 0.25 | 0.0344 | 0.128 | 0.12 | 0.12 | 0.13 | 0.13 | 0.12 | 0.12712 | 0.12826 | 0.14 | 0.119 | 0.139
| Mean |  | 0.12883 | 0.26167 | 0.0344 | 0.1275 | 0.125 | 0.12 | 0.12667 | 0.13 | 0.12167 | 0.12621 | 0.12705 | 0.13833 | 0.1245 | 0.135
| Median |  | 0.129 | 0.25 | 0.03555 | 0.1275 | 0.125 | 0.12 | 0.13 | 0.13 | 0.12 | 0.12609 | 0.12711 | 0.14 | 0.121 | 0.136
| Std Dev. |  | 0.001169 | 0.027869 | 0.0029093 | 0.00054772 | 0.010488 | 0 | 0.005164 | 0 | 0.0040825 | 0.0005386 | 0.0018589 | 0.0040825 | 0.0079687 | 0.0046043
| Rel.Std.Dev. |  | 0.9074% | 10.65% | 8.457% | 0.4296% | 8.39% | 0% | 4.077% | 0% | 3.355% | 0.4268% | 1.463% | 2.951% | 6.401% | 3.411%
| PDM3 |  | 1.777% | 106.7% | -72.82% | 0.7237% | -1.251% | -5.201% | 0.06539% | 2.699% | -3.885% | -0.2984% | 0.3683% | 9.282% | -1.646% | 6.649%
| Z-Score (Absolute) |  | 0.48 | 33.06 | 22.68 | 0.16 | 0.46 | 1.68 | 0.05 | 0.77 | 1.27 | 0.16 | 0.05 | 2.81 | 0.58 | 2
| Analytical results for Pb in OREAS 143 (Indicative Value 85 ppm) | [15 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 14 | 15 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 10 | < 50 | 60 | < 50 | < 50 | < 50 | 140 | 2275.9068 | 190 |  |  |  |  | 1
| 1 | 2 | < 100 | < 10 | < 50 | < 30 | < 50 | < 50 | < 50 | 110 | 2360.1135 | 190 |  |  |  |  | 6
| 1 | 3 | < 100 | < 10 | < 50 | 30 | < 50 | 160 | < 50 | 130 | 2350.4683 | 190 |  |  |  |  | 16
| 1 | 4 | < 100 | < 10 | < 50 | < 30 | < 50 | < 50 | < 50 | 120 | 2282.8457 | 190 |  |  |  |  | 84.5486
| 1 | 5 | < 100 | < 10 | < 50 | < 30 | < 50 | < 50 | < 50 | 130 | 2412.1852 | 190 |  |  |  |  | 12
| 1 | 6 | < 100 | < 10 | < 50 | < 30 | < 50 | < 50 | < 50 | 120 | 2305.8049 | 90
| Mean |  |  |  |  | 45 |  | 160 |  | 125 | 2331.2207 | 173.3333
| Median |  |  |  |  | 45 |  | 160 |  | 125 | 2328.1366 | 190
| Std Dev. |  |  |  |  | 21.2132 |  |  |  | 10.4881 | 52.5476 | 40.8248
| Rel.Std.Dev. |  |  |  |  | 47.14% |  |  |  | 8.39% | 2.254% | 23.55%
| PDM3 |  |  |  |  | -46.78% |  | 89.24% |  | 47.84% | 2657% | 105%
| Z-Score (Absolute) |  | 0.74 | 1.69 | 0.61 | 0.61 | 0.61 | 0.61 | 0.61 | 4.79 | 123.8 | 7.23
| Analytical results for Rb in OREAS 143 (Indicative Value 108 ppm) | [15 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 06 | 07 | 08 | 09 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | 120 | 110 | 90 | 110 | [9 empty] | 1
| 1 | 2 | < 100 | 110 | 140 | 90 | 110 | [9 empty] | 7
| 1 | 3 | < 100 | 120 | 170 | 90 | 110 | [9 empty] | 16
| 1 | 4 | < 100 | 110 | 130 | 100 | 120 | [9 empty] | 108.25
| 1 | 5 | < 100 | 120 | 100 | 90 | 90 | [9 empty] | 13
| 1 | 6 | < 100 | 110 | 110 | 90 | 110
| Mean |  |  | 115 | 126.6667 | 91.6667 | 108.3333
| Median |  |  | 115 | 120 | 90 | 110
| Std Dev. |  |  | 5.4772 | 25.8199 | 4.0825 | 9.8319
| Rel.Std.Dev. |  |  | 4.763% | 20.38% | 4.454% | 9.076%
| PDM3 |  |  | 6.236% | 17.01% | -15.32% | 0.07698%
| Z-Score (Absolute) |  | 2.36 | 0.27 | 0.74 | 0.67 | 0
| Analytical results for Sb in OREAS 143 (Indicative Value < 50 ppm) | [15 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 09 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 10 | < 50 | 60 | < 50 | < 50 | [8 empty] | 1
| 1 | 2 | < 100 | < 10 | < 50 | 60 | < 50 | < 50 | [8 empty] | 8
| 1 | 3 | < 100 | < 10 | < 50 | < 50 | < 50 | < 50 | [8 empty] | 16
| 1 | 4 | < 100 | < 10 | < 50 | < 50 | < 50 | < 50 | [8 empty] | < 50
| 1 | 5 | < 100 | < 10 | 50 | 90 | < 50 | < 50 | [8 empty] | 14
| 1 | 6 | < 100 | < 10 | 60 | 50 | < 50 | < 50
| Mean |  |  |  | 55 | 65
| Median |  |  |  | 55 | 60
| Std Dev. |  |  |  | 7.0711 | 17.3205
| Rel.Std.Dev. |  |  |  | 12.86% | 26.65%
PDM3
| Z-Score (Absolute) |  | 1.08 | 1.35 | 0.27 | 1.17 | 0.27 | 0.27
| Analytical results for Sc in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 02 | 03 | [12 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | 12 | [12 empty] | 1
| 1 | 2 | < 100 | < 10 | [12 empty] | 1
| 1 | 3 | < 100 | < 10 | [12 empty] | 16
| 1 | 4 | < 100 | < 10 | [12 empty] | < 100
| 1 | 5 | < 100 | < 10 | [12 empty] | 7
| 1 | 6 | < 100 | < 10
| Mean |  |  | 12
| Median |  |  | 12
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Se in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 22
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 143 (Certified Value 54.17 wt.%) | [15 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 53.71 | 50.12 | 53.394 | 53.92 | 54.12 | 54.2 | 54.2 | 54.7 | 54 | 52.692 | 56.3904 | 54.59 | 53.94 | 55.149 | 1
| 1 | 2 | 53.88 | 50.23 | 53.48 | 54.02 | 54.4 | 54.7 | 54.1 | 55.1 | 54.2 | 52.8324 | 56.3956 | 54.85 | 57.28 | 54.374 | #N/A
| 1 | 3 | 53.65 | 50.14 | 53.523 | 53.81 | 54.06 | 54.2 | 54.5 | 54.9 | 54.2 | 52.5571 | 55.6448 | 54.83 | 49.67 | 54.021 | 16
| 1 | 4 | 53.73 | 50.25 | 53.245 | 53.77 | 54.7 | 54.2 | 54.3 | 55.6 | 54 | 52.6942 | 55.5835 | 55.07 | 53.25 | 55.209 | 54.1658
| 1 | 5 | 53.68 | 49.96 | 52.86 | 53.83 | 53.74 | 54.9 | 54.9 | 54.8 | 54.3 | 52.6788 | 56.0583 | 55.43 | 56.03 | 54.616 | 21
| 1 | 6 | 53.67 | 49.98 | 53.245 | 53.81 | 54.29 | 54.6 | 54.6 | 54.9 | 53.9 | 52.6853 | 55.6828 | 55.29 | 55.09 | 54.369
| Mean |  | 53.72 | 50.1133 | 53.2912 | 53.86 | 54.2183 | 54.4667 | 54.4333 | 55 | 54.1 | 52.69 | 55.9592 | 55.01 | 54.21 | 54.623
| Median |  | 53.695 | 50.13 | 53.3195 | 53.82 | 54.205 | 54.4 | 54.4 | 54.9 | 54.1 | 52.6886 | 55.8706 | 54.96 | 54.515 | 54.495
| Std Dev. |  | 0.083427 | 0.12193 | 0.24089 | 0.092952 | 0.32671 | 0.30768 | 0.29439 | 0.32249 | 0.15492 | 0.087305 | 0.37487 | 0.31394 | 2.6502 | 0.47097
| Rel.Std.Dev. |  | 0.1553% | 0.2433% | 0.452% | 0.1726% | 0.6026% | 0.5649% | 0.5408% | 0.5863% | 0.2864% | 0.1657% | 0.6699% | 0.5707% | 4.889% | 0.8622%
| PDM3 |  | -0.8231% | -7.482% | -1.615% | -0.5646% | 0.09695% | 0.5554% | 0.4939% | 1.54% | -0.1215% | -2.725% | 3.311% | 1.559% | 0.08157% | 0.844%
| Z-Score (Absolute) |  | 0.74 | 6.12 | 1.38 | 0.53 | 0.01 | 0.38 | 0.33 | 1.17 | 0.17 | 2.28 | 2.61 | 1.19 | 0.01 | 0.61
| Analytical results for Sn in OREAS 143 (Certified Value 0.252 wt.%) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 14 | 15 | 17 | 18 |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.244 | 0.28 | 0.26 | 0.242 | 0.26 | 0.255 | 0.248 | 0.25 | 0.27 | 0.25065 | 0.252 | 0.254 | 0.2324 |  | 1
| 1 | 2 | 0.244 | 0.27 | 0.25 | 0.25 | 0.25 | 0.256 | 0.242 | 0.256 | 0.261 | 0.25343 | 0.252 | 0.268 | 0.2347 |  | 20
| 1 | 3 | 0.244 | 0.24 | 0.25 | 0.244 | 0.24 | 0.256 | 0.252 | 0.254 | 0.259 | 0.25503 | 0.26 | 0.244 | 0.2426 |  | 16
| 1 | 4 | 0.244 | 0.28 | 0.26 | 0.256 | 0.25 | 0.259 | 0.253 | 0.254 | 0.262 | 0.25275 | 0.244 | 0.244 | 0.2347 |  | 0.25212
| 1 | 5 | 0.244 | 0.27 | 0.25 | 0.256 | 0.25 | 0.253 | 0.248 | 0.248 | 0.263 | 0.25955 | 0.252 | 0.254 | 0.2292 |  | 22
| 1 | 6 | 0.244 | 0.25 | 0.26 | 0.247 | 0.24 | 0.258 | 0.246 | 0.253 | 0.252 | 0.25395 | 0.244 | 0.258 | 0.23
| Mean |  | 0.244 | 0.265 | 0.255 | 0.24917 | 0.24833 | 0.25617 | 0.24817 | 0.2525 | 0.26117 | 0.25423 | 0.25067 | 0.25367 | 0.23393
| Median |  | 0.244 | 0.27 | 0.255 | 0.2485 | 0.25 | 0.256 | 0.248 | 0.2535 | 0.2615 | 0.25369 | 0.252 | 0.254 | 0.23355
| Std Dev. |  | 0 | 0.016432 | 0.0054772 | 0.005947 | 0.0075277 | 0.002137 | 0.0040208 | 0.0029496 | 0.0058452 | 0.0029908 | 0.0060222 | 0.0090701 | 0.0048273
| Rel.Std.Dev. |  | 0% | 6.201% | 2.148% | 2.387% | 3.031% | 0.8342% | 1.62% | 1.168% | 2.238% | 1.176% | 2.402% | 3.576% | 2.064%
| PDM3 |  | -3.22% | 5.109% | 1.143% | -1.171% | -1.501% | 1.606% | -1.567% | 0.1514% | 3.589% | 0.8363% | -0.5757% | 0.6142% | -7.213%
| Z-Score (Absolute) |  | 1.53 | 2.3 | 0.46 | 0.61 | 0.77 | 0.67 | 0.8 | 0 | 1.59 | 0.32 | 0.32 | 0.21 | 3.41
Results from laboratory 2 were removed due to their high Z-score and unsatisfactory performance on the control sample.
| Analytical results for SO3 in OREAS 143 (Certified Value 3.08 wt.%) | [15 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 12 | 15 | 17 | 18 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 3.098 | 3.41 | 2.9267 | 2.6245 | 3.121 | 3.097 | 3.097 | 3.047 | 3.0536 | 3.146 | 3.003 | 2.813 |  |  | 1
| 1 | 2 | 3.036 | 3.49 | 2.8942 | 2.5721 | 3.097 | 3.097 | 3.146 | 3.047 | 3.0311 | 3.146 | 3.204 | 2.68 |  |  | #N/A
| 1 | 3 | 3.115 | 3.45 | 2.8942 | 2.6545 | 3.146 | 3.146 | 3.146 | 3.072 | 3.1271 | 3.171 | 2.794 | 2.613 |  |  | 16
| 1 | 4 | 3.101 | 3.47 | 2.8892 | 2.5696 | 3.047 | 3.097 | 3.146 | 3.047 | 3.0042 | 3.146 | 3.002 | 2.787 |  |  | 3.0759
| 1 | 5 | 3.081 | 3.46 | 2.8643 | 2.6146 | 3.146 | 3.097 | 3.121 | 3.047 | 3.0325 | 3.171 | 3.135 | 2.666 |  |  | 23
| 1 | 6 | 3.122 | 3.48 | 2.8967 | 2.642 | 3.097 | 3.097 | 3.146 | 3.047 | 3.0195 | 3.171 | 3.127 | 2.713
| Mean |  | 3.0922 | 3.46 | 2.8942 | 2.6129 | 3.109 | 3.1052 | 3.1337 | 3.0512 | 3.0447 | 3.1585 | 3.0442 | 2.712
| Median |  | 3.0995 | 3.465 | 2.8942 | 2.6195 | 3.109 | 3.097 | 3.146 | 3.047 | 3.0318 | 3.1585 | 3.065 | 2.6965
| Std Dev. |  | 0.030993 | 0.028284 | 0.019907 | 0.035368 | 0.037454 | 0.020004 | 0.020559 | 0.010206 | 0.04355 | 0.013693 | 0.14609 | 0.075852
| Rel.Std.Dev. |  | 1.002% | 0.8175% | 0.6878% | 1.354% | 1.205% | 0.6442% | 0.6561% | 0.3345% | 1.43% | 0.4335% | 4.799% | 2.797%
| PDM3 |  | 0.53% | 12.49% | -5.906% | -15.05% | 1.077% | 0.9527% | 1.879% | -0.8029% | -1.014% | 2.687% | -1.031% | -11.83%
| Z-Score (Absolute) |  | 0.28 | 5.24 | 2.39 | 6.18 | 0.51 | 0.45 | 0.84 | 0.28 | 0.36 | 1.18 | 0.37 | 4.84
| Analytical results for SrO in OREAS 143 (Certified Value 120 ppm) | [15 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 09 | 12 | 14 | 17 | 18 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 300 | 130 | 120 | 120 | 120 | 120 | 155.6453 | < 50 | 200 | 90 |  |  |  |  | 1
| 1 | 2 | 300 | 142 | 120 | 120 | 120 | 120 | 155.7163 | < 50 | 210 | 80 |  |  |  |  | #N/A
| 1 | 3 | 300 | 154 | 120 | 120 | 120 | 120 | 156.3128 | < 50 | 120 | 70 |  |  |  |  | 16
| 1 | 4 | 300 | 154 | 120 | 120 | 120 | 120 | 154.6777 | < 50 | 230 | 50 |  |  |  |  | 120.1151
| 1 | 5 | 300 | 154 | 120 | 120 | 120 | 120 | 154.4774 | < 50 | 240 | 60 |  |  |  |  | 24
| 1 | 6 | 300 | 142 | 120 | 120 | 120 | 120 | 154.6401 | < 50 | 240 | 50
| Mean |  | 300 | 146 | 120 | 120 | 120 | 120 | 155.2449 |  | 206.6667 | 66.6667
| Median |  | 300 | 148 | 120 | 120 | 120 | 120 | 155.1615 |  | 220 | 65
| Std Dev. |  | 0 | 9.798 | 0 | 0 | 0 | 0 | 0.74826 |  | 45.4606 | 16.3299
| Rel.Std.Dev. |  | 0% | 6.711% | 0% | 0% | 0% | 0% | 0.482% |  | 22% | 24.49%
| PDM3 |  | 149.8% | 21.55% | -0.09581% | -0.09581% | -0.09581% | -0.09581% | 29.25% |  | 72.06% | -44.5%
| Z-Score (Absolute) |  | 3.8 | 0.58 | 0 | 0 | 0 | 0 | 0.77 | 1.95 | 1.85 | 1.08
| Analytical results for Ta in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Ta | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 3
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 9
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Te | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 4
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 10
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 143 (Certified Value 0.576 wt.%) | [15 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.58 | 0.79 | 0.5621 | 0.57 | 0.58 | 0.57 | 0.56 | 0.57 | 0.59 | 0.56683 | 0.6101 | 0.59 | 0.61 | 0.582 | 1
| 1 | 2 | 0.58 | 0.72 | 0.5655 | 0.58 | 0.57 | 0.57 | 0.56 | 0.58 | 0.57 | 0.56727 | 0.60943 | 0.59 | 0.64 | 0.57 | #N/A
| 1 | 3 | 0.58 | 0.72 | 0.5605 | 0.58 | 0.57 | 0.58 | 0.56 | 0.58 | 0.58 | 0.57496 | 0.6084 | 0.59 | 0.56 | 0.564 | 16
| 1 | 4 | 0.59 | 0.75 | 0.5271 | 0.57 | 0.57 | 0.56 | 0.56 | 0.6 | 0.58 | 0.56479 | 0.63641 | 0.58 | 0.61 | 0.577 | 0.57564
| 1 | 5 | 0.58 | 0.79 | 0.5655 | 0.57 | 0.56 | 0.57 | 0.58 | 0.57 | 0.58 | 0.57847 | 0.61241 | 0.59 | 0.64 | 0.57 | 25
| 1 | 6 | 0.59 | 0.76 | 0.5305 | 0.56 | 0.58 | 0.6 | 0.57 | 0.58 | 0.58 | 0.55703 | 0.61083 | 0.59 | 0.62 | 0.576
| Mean |  | 0.58333 | 0.755 | 0.55187 | 0.57167 | 0.57167 | 0.575 | 0.565 | 0.58 | 0.58 | 0.56822 | 0.6146 | 0.58833 | 0.61333 | 0.57317
| Median |  | 0.58 | 0.755 | 0.5613 | 0.57 | 0.57 | 0.57 | 0.56 | 0.58 | 0.58 | 0.56705 | 0.61047 | 0.59 | 0.615 | 0.573
| Std Dev. |  | 0.005164 | 0.031464 | 0.018005 | 0.0075277 | 0.0075277 | 0.013784 | 0.0083666 | 0.010954 | 0.0063246 | 0.0076197 | 0.010769 | 0.0040825 | 0.029439 | 0.0064005
| Rel.Std.Dev. |  | 0.8853% | 4.167% | 3.263% | 1.317% | 1.317% | 2.397% | 1.481% | 1.889% | 1.09% | 1.341% | 1.752% | 0.6939% | 4.8% | 1.117%
| PDM3 |  | 1.337% | 31.16% | -4.13% | -0.6901% | -0.6901% | -0.111% | -1.848% | 0.7576% | 0.7576% | -1.288% | 6.768% | 2.205% | 6.548% | -0.4295%
| Z-Score (Absolute) |  | 0.52 | 15.84 | 2.29 | 0.52 | 0.52 | 0.22 | 1.12 | 0.22 | 0.22 | 0.83 | 3.31 | 0.97 | 3.2 | 0.39
Results from laboratory 3 were removed due to their high Z-score and unsatisfactory performance on the control sample.
| Analytical results for Tl in OREAS 143 (Indicative Value < 100 ppm) | [15 empty] | Indicative
Tl | Round | Replicate | Lab | [13 empty] | 1
| No | No | 02 | [13 empty] | ppm
|  |  | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | [13 empty] | 1
| 1 | 2 | < 100 | [13 empty] | 5
| 1 | 3 | < 100 | [13 empty] | 16
| 1 | 4 | < 100 | [13 empty] | < 100
| 1 | 5 | < 100 | [13 empty] | 11
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 143 (Certified Value 135 ppm) | [15 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 180 | 125 | 180 | 110 | < 180 | 180 | 180 | 180 | 135.2996 | < 200 | < 100 | < 50 | 120 |  | 1
| 1 | 2 | < 180 | 125 | 170 | 100 | < 180 | 180 | 180 | 180 | 131.6578 | < 200 | 100 | 80 | 150 |  | #N/A
| 1 | 3 | < 180 | 125 | 170 | 120 | < 180 | 180 | 180 | 180 | 134.2284 | < 200 | 100 | < 50 | 130 |  | 16
| 1 | 4 | < 180 | 125 | 180 | 90 | < 180 | 180 | 180 | 180 | 134.4784 | < 200 | 100 | < 50 | 120 |  | 135.2579
| 1 | 5 | < 180 | 125 | 180 | 90 | < 180 | 180 | 180 | 180 | 137.656 | < 200 | 100 | < 50 | 140 |  | 26
| 1 | 6 | < 180 | 142.8 | 180 | 130 | < 180 | 180 | 180 | 180 | 136.7813 | < 200 | < 100 | < 50 | 120
| Mean |  |  | 127.9667 | 176.6667 | 106.6667 |  | 180 | 180 | 180 | 135.0169 |  | 100 | 80 | 130
| Median |  |  | 125 | 180 | 105 |  | 180 | 180 | 180 | 134.889 |  | 100 | 80 | 125
| Std Dev. |  |  | 7.2668 | 5.164 | 16.3299 |  | 0 | 0 | 0 | 2.1144 |  | 0 |  | 12.6491
| Rel.Std.Dev. |  |  | 5.679% | 2.923% | 15.31% |  | 0% | 0% | 0% | 1.566% |  | 0% |  | 9.73%
| PDM3 |  |  | -5.391% | 30.61% | -21.14% |  | 33.08% | 33.08% | 33.08% | -0.1782% |  | -26.07% | -40.85% | -3.887%
| Z-Score (Absolute) |  | 1.08 | 0.41 | 2.29 | 0.41 | NA | NA | NA | NA | 0.68 | 0.67 | NA | 3.2 | 0.49
Results from laboratories 6, 7, 8, 9 and 15 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 143 (Indicative Value 15.4 ppm) | [15 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 09 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [8 empty] | 1
1
| 1 | 1 | < 100 | < 10 | 20 | < 10 | 10 | 20 | [8 empty] | 1
| 1 | 2 | < 100 | < 10 | 20 | < 10 | 10 | 20 | [8 empty] | 6
| 1 | 3 | < 100 | < 10 | 30 | 60 | 20 | 20 | [8 empty] | 16
| 1 | 4 | < 100 | < 10 | 20 | < 10 | 20 | 10 | [8 empty] | 15.4167
| 1 | 5 | < 100 | < 10 | 30 | < 10 | 10 | 20 | [8 empty] | 12
| 1 | 6 | < 100 | < 10 | 30 | < 10 | 10 | 20
| Mean |  |  |  | 25 | 60 | 13.3333 | 18.3333
| Median |  |  |  | 25 | 60 | 10 | 20
| Std Dev. |  |  |  | 5.4772 |  | 5.164 | 4.0825
| Rel.Std.Dev. |  |  |  | 21.91% |  | 38.73% | 22.27%
| PDM3 |  |  |  | 62.16% | 289.2% | -13.51% | 18.92%
| Z-Score (Absolute) |  | 3.9 | 1.3 | 1.01 | 0.24 | 0.34 | 0.24
| Analytical results for Y in OREAS 143 (Indicative Value 47.6 ppm) | [15 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 06 | 07 | 08 | 09 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [10 empty] | 1
1
| 1 | 1 | < 39 | 47 | < 39 | 55 | [10 empty] | 1
| 1 | 2 | < 39 | 63 | < 39 | 55 | [10 empty] | 7
| 1 | 3 | 39 | 63 | < 39 | 55 | [10 empty] | 16
| 1 | 4 | < 39 | 71 | < 39 | 63 | [10 empty] | 47.6415
| 1 | 5 | 39 | 126 | < 39 | 47 | [10 empty] | 13
| 1 | 6 | 39 | 47 | < 39 | 55
| Mean |  | 39 | 69.5 |  | 55
| Median |  | 39 | 63 |  | 55
| Std Dev. |  | 0 | 29.2968 |  | 5.0596
| Rel.Std.Dev. |  | 0% | 42.15% |  | 9.199%
| PDM3 |  | -18.14% | 45.88% |  | 15.45%
| Z-Score (Absolute) |  | 0.49 | 1.04 | 0.86 | 0.49
| Analytical results for Zn in OREAS 143 (Certified Value 108 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 14 | 15 | 18 |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 200 | < 50 | 110 | 60 | 110 | 100 | 90 | 110 | < 160 | 160 | 112 |  |  |  | 1
| 1 | 2 | 100 | < 50 | 110 | 90 | 120 | 80 | 80 | 120 | < 160 | 160 | 112 |  |  |  | #N/A
| 1 | 3 | 200 | < 50 | 110 | 80 | 110 | 120 | 90 | 100 | < 160 | 160 | 104 |  |  |  | 16
| 1 | 4 | 100 | < 50 | 120 | 60 | 110 | 90 | 90 | 110 | < 160 | 80 | 104 |  |  |  | 108.4297
| 1 | 5 | 100 | < 50 | 110 | 60 | 110 | 100 | 80 | 110 | < 160 | 80 | 112 |  |  |  | 27
| 1 | 6 | 300 | < 50 | 110 | 60 | 110 | 100 | 90 | 100 | < 160 | 160 | 104
| Mean |  | 166.6667 |  | 111.6667 | 68.3333 | 111.6667 | 98.3333 | 86.6667 | 108.3333 |  | 133.3333 | 108
| Median |  | 150 |  | 110 | 60 | 110 | 100 | 90 | 110 |  | 160 | 108
| Std Dev. |  | 81.6497 |  | 4.0825 | 13.2916 | 4.0825 | 13.2916 | 5.164 | 7.5277 |  | 41.3118 | 4.3818
| Rel.Std.Dev. |  | 48.99% |  | 3.656% | 19.45% | 3.656% | 13.52% | 5.958% | 6.949% |  | 30.98% | 4.057%
| PDM3 |  | 53.71% |  | 2.985% | -36.98% | 2.985% | -9.311% | -20.07% | -0.08891% |  | 22.97% | -0.3963%
| Z-Score (Absolute) |  | 1.82 | 2.59 | 0.1 | 1.24 | 0.1 | 0.31 | 0.67 | 0 | 0.87 | 0.8 | 0
| Analytical results for ZrO2 in OREAS 143 (Indicative Value 271 ppm) | [15 empty] | Indicative
ZrO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 06 | 07 | 08 | 09 | 14 | 18 | [6 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 270 | 54 | 270 | 410 | 270 | 140 | < 500 | 250 | [6 empty] | 1
| 1 | 2 | 270 | 41 | 270 | 270 | 270 | 140 | < 500 | 290 | [6 empty] | 8
| 1 | 3 | 270 | 27 | 270 | 270 | 270 | 140 | < 500 | 260 | [6 empty] | 16
| 1 | 4 | 270 | 27 | 270 | 140 | 270 | 140 | < 500 | 280 | [6 empty] | 270.9479
| 1 | 5 | 270 | 27 | 270 | 270 | 270 | 140 | < 500 | 270 | [6 empty] | 14
| 1 | 6 | 270 | 27 | 270 | 140 | 270 | 140 | < 500 | 290
| Mean |  | 270 | 33.8333 | 270 | 250 | 270 | 140 |  | 273.3333
| Median |  | 270 | 27 | 270 | 270 | 270 | 140 |  | 275
| Std Dev. |  | 0 | 11.3563 | 0 | 100.995 | 0 | 0 |  | 16.3299
| Rel.Std.Dev. |  | 0% | 33.57% | 0% | 40.4% | 0% | 0% |  | 5.974%
| PDM3 |  | -0.3498% | -87.51% | -0.3498% | -7.731% | -0.3498% | -48.33% |  | 0.8804%
| Z-Score (Absolute) |  | 0.6 | 13.56 | 0.6 | 0.74 | 0.6 | 7.49 | 0.6 | 0.79
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C761: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 143 (Indicative Value 0.295 wt.%) | [15 empty] | Indicative
H2O- | Round | Replicate | Lab | [13 empty] | 1
| No | No | 15 | [13 empty] | wt.%
|  |  | LOI*TGA | [13 empty] | 3
3
| 1 | 1 | 0.3 | [13 empty] | 1
| 1 | 2 | 0.3 | [13 empty] | 10
| 1 | 3 | 0.31 | [13 empty] | 16
| 1 | 4 | 0.29 | [13 empty] | 0.295
| 1 | 5 | 0.28 | [13 empty] | 16
| 1 | 6 | 0.29
| Mean |  | 0.295
| Median |  | 0.295
| Std Dev. |  | 0.010488
| Rel.Std.Dev. |  | 3.555%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 143 (Certified Value 6.66 wt.%) | [15 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 17 | 18 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | 2
3
| 1 | 1 | 6.72 | 6.63 | 6.79 | 6.4682 | 6.8 | 6.65 | 5.78 | 6.37 | 6.81 | 6.684 | 6.4272 | 6.78 | 6.62 | 4.12 | 1
| 1 | 2 | 6.78 | 6.67 | 6.8 | 6.4261 | 6.73 | 6.59 | 6 | 6.3 | 6.79 | 6.7632 | 6.5326 | 6.8 | 6.52 | 5.75 | #N/A
| 1 | 3 | 6.74 | 6.61 | 6.74 | 6.4249 | 6.78 | 6.62 | 5.76 | 6.2 | 6.82 | 6.8873 | 6.225 | 6.85 | 6.41 | 6.9 | 16
| 1 | 4 | 6.71 | 6.63 | 6.74 | 6.3593 | 6.7 | 6.68 | 5.8 | 6.28 | 6.86 | 6.8473 | 6.4988 | 6.84 | 6.74 | 4.15 | 6.6584
| 1 | 5 | 6.78 | 6.66 | 6.63 | 6.4778 | 6.68 | 6.58 | 5.8 | 6.37 | 6.91 | 6.994 | 6.5322 | 6.89 | 6.68 | 5.29 | 29
| 1 | 6 | 6.81 | 6.68 | 6.65 | 6.418 | 6.7 | 6.65 | 5.81 | 6.53 | 6.98 | 6.9489 | 6.6236 | 6.86 | 6.43 | 5.75
| Mean |  | 6.7567 | 6.6467 | 6.725 | 6.429 | 6.7317 | 6.6283 | 5.825 | 6.3417 | 6.8617 | 6.8541 | 6.4732 | 6.8367 | 6.5667 | 5.3267
| Median |  | 6.76 | 6.645 | 6.74 | 6.4255 | 6.715 | 6.635 | 5.8 | 6.335 | 6.84 | 6.8673 | 6.5155 | 6.845 | 6.57 | 5.52
| Std Dev. |  | 0.039328 | 0.027325 | 0.07064 | 0.042245 | 0.048339 | 0.038687 | 0.087579 | 0.11197 | 0.071949 | 0.11569 | 0.13709 | 0.040332 | 0.13501 | 1.0654
| Rel.Std.Dev. |  | 0.5821% | 0.4111% | 1.05% | 0.6571% | 0.7181% | 0.5837% | 1.503% | 1.766% | 1.049% | 1.688% | 2.118% | 0.5899% | 2.056% | 20%
| PDM3 |  | 1.476% | -0.1765% | 1% | -3.445% | 1.1% | -0.4518% | -12.52% | -4.757% | 3.053% | 2.939% | -2.781% | 2.677% | -1.378% | -20%
| Z-Score (Absolute) |  | 0.44 | 0.03 | 0.32 | 0.77 | 0.35 | 0.03 | 3.01 | 1.1 | 0.83 | 0.8 | 0.61 | 0.74 | 0.26 | 4.86
some laboratories include LOI in their fusion package.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 143 (Indicative Value < 0.5 ppm) | [13 empty] | Indicative
Ag | Round | Replicate | Lab | [11 empty] | 1
| No | No | 02 | [11 empty] | ppm
|  |  | PF*OES | [11 empty] | 3
3
| 1 | 1 | < 0.5 | [11 empty] | 1
| 1 | 2 | < 0.5 | [11 empty] | 12
| 1 | 3 | < 0.5 | [11 empty] | 16
| 1 | 4 | < 0.5 | [11 empty] | < 0.5
| 1 | 5 | < 0.5 | [11 empty] | 18
| 1 | 6 | < 0.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Al in OREAS 143 (Certified Value 6.51 wt.%) | [13 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS |  |  | 2
3
| 1 | 1 | 6.72 | 2.04 | 6.35 | 6.53 | 6.51 | 6.642 | 6.483 | 6.7839 | 6.506 | 6.4696 |  |  | 1
| 1 | 2 | 6.71 | 2.03 | 6.36 | 6.58 | 6.351 | 6.88 | 6.589 | 6.7415 | 6.084 | 6.4363 |  |  | #N/A
| 1 | 3 | 6.69 | 2.06 | 6.39 | 6.69 | 6.298 | 6.668 | 6.483 | 6.757 | 6.65 | 6.4321 |  |  | 16
| 1 | 4 | 6.82 | 2.07 | 6.28 | 6.27 | 6.351 | 6.668 | 6.457 | 6.7103 | 6.306 | 6.4662 |  |  | 6.515
| 1 | 5 | 6.71 | 2.05 | 6.32 | 6.53 | 6.43 | 6.616 | 6.668 | 6.7854 | 6.317 | 6.4556 |  |  | 31
| 1 | 6 | 6.57 | 2.09 | 6.22 | 6.15 | 6.351 | 6.43 | 6.642 | 6.7847 | 6.323 | 6.3919
| Mean |  | 6.7033 | 2.0567 | 6.32 | 6.4583 | 6.3818 | 6.6507 | 6.5537 | 6.7605 | 6.3643 | 6.442
| Median |  | 6.71 | 2.055 | 6.335 | 6.53 | 6.351 | 6.655 | 6.536 | 6.7705 | 6.32 | 6.446
| Std Dev. |  | 0.079917 | 0.021602 | 0.061644 | 0.20459 | 0.075666 | 0.14364 | 0.091055 | 0.030486 | 0.19371 | 0.028895
| Rel.Std.Dev. |  | 1.192% | 1.05% | 0.9754% | 3.168% | 1.186% | 2.16% | 1.389% | 0.4509% | 3.044% | 0.4485%
| PDM3 |  | 2.891% | -68.43% | -2.993% | -0.8693% | -2.044% | 2.083% | 0.594% | 3.768% | -2.312% | -1.121%
| Z-Score (Absolute) |  | 1.46 | 25.35 | 0.75 | 0.05 | 0.39 | 1.16 | 0.6 | 1.79 | 0.5 | 0.05
| Analytical results for As in OREAS 143 (Certified Value 0.115 wt.%) | [13 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 11 | 12 | 13 | 17 |  | wt.%
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.1194 | 0.10201 | 0.114 | 0.117 | 0.12 | 0.11 | 0.12 | 0.114 | 0.1027 | 0.112 | 0.10996 |  | 1
| 1 | 2 | 0.1212 | 0.10182 | 0.115 | 0.114 | 0.12 | 0.12 | 0.13 | 0.119 | 0.10342 | 0.1133 | 0.11348 |  | #N/A
| 1 | 3 | 0.1211 | 0.10273 | 0.114 | 0.117 | 0.12 | 0.12 | 0.12 | 0.111 | 0.10663 | 0.1105 | 0.11343 |  | 16
| 1 | 4 | 0.1196 | 0.10265 | 0.114 | 0.115 | 0.12 | 0.11 | 0.12 | 0.119 | 0.10389 | 0.1113 | 0.11168 |  | 0.11533
| 1 | 5 | 0.1199 | 0.10312 | 0.116 | 0.116 | 0.12 | 0.12 | 0.13 | 0.111 | 0.1059 | 0.1106 | 0.113 |  | 32
| 1 | 6 | 0.1197 | 0.10304 | 0.114 | 0.116 | 0.12 | 0.11 | 0.12 | 0.118 | 0.10435 | 0.1123 | 0.1163
| Mean |  | 0.12015 | 0.10256 | 0.1145 | 0.11583 | 0.12 | 0.115 | 0.12333 | 0.11533 | 0.10448 | 0.11167 | 0.11297
| Median |  | 0.1198 | 0.10269 | 0.114 | 0.116 | 0.12 | 0.115 | 0.12 | 0.116 | 0.10412 | 0.11165 | 0.11321
| Std Dev. |  | 0.00079183 | 0.00053499 | 0.00083666 | 0.001169 | 0 | 0.0054772 | 0.005164 | 0.0038297 | 0.0015036 | 0.0010783 | 0.0021112
| Rel.Std.Dev. |  | 0.659% | 0.5216% | 0.7307% | 1.009% | 0% | 4.763% | 4.187% | 3.321% | 1.439% | 0.9656% | 1.869%
| PDM3 |  | 4.182% | -11.07% | -0.7174% | 0.4388% | 4.052% | -0.2838% | 6.942% | 0.005224% | -9.404% | -3.174% | -2.04%
| Z-Score (Absolute) |  | 1.04 | 2.52 | 0.1 | 0.17 | 1.01 | 0 | 1.69 | 0.07 | 2.13 | 0.67 | 0.41
| Analytical results for B in OREAS 143 (Indicative Value 664 ppm) | [13 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 04 | 05 | 17 | [8 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | [8 empty] | 0
0
| 1 | 1 | 709 | 640 | 600 | 700 | [8 empty] | 1
| 1 | 2 | 716 | 650 | 560 | 737 | [8 empty] | 13
| 1 | 3 | 705 | 650 | 610 | 721 | [8 empty] | 16
| 1 | 4 | 756 | 640 | 610 | 683 | [8 empty] | 663.9
| 1 | 5 | 706 | 650 | 590 | 694 | [8 empty] | 19
| 1 | 6 | 717 | 640 | 580 | 715
| Mean |  | 718.1667 | 645 | 591.6667 | 708.3333
| Median |  | 712.5 | 645 | 595 | 707.5
| Std Dev. |  | 19.1981 | 5.4772 | 19.4079 | 19.7146
| Rel.Std.Dev. |  | 2.673% | 0.8492% | 3.28% | 2.783%
| PDM3 |  | 8.174% | -2.847% | -10.88% | 6.693%
| Z-Score (Absolute) |  | 0.76 | 0.58 | 1.57 | 0.58
| Analytical results for Ba in OREAS 143 (Certified Value 649 ppm) | [13 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 13 | 14 | [5 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 669 | 197.4 | 657 | 656 | 652.3 | 607 | 641.3421 | [5 empty] | 1
| 1 | 2 | 683 | 199 | 664 | 660 | 654.67 | 585 | 648.8184 | [5 empty] | #N/A
| 1 | 3 | 679 | 202.9 | 661 | 655 | 646.26 | 617 | 642.6355 | [5 empty] | 16
| 1 | 4 | 663 | 199.8 | 669 | 648 | 650.04 | 605 | 640.0061 | [5 empty] | 649.4503
| 1 | 5 | 665 | 200 | 659 | 678 | 644.52 | 614 | 646.3952 | [5 empty] | 33
| 1 | 6 | 656 | 196.9 | 663 | 672 | 648.24 | 606 | 649.1839
| Mean |  | 669.1667 | 199.3333 | 662.1667 | 661.5 | 649.3383 | 605.6667 | 644.7302
| Median |  | 667 | 199.4 | 662 | 658 | 649.14 | 606.5 | 644.5154
| Std Dev. |  | 10.1669 | 2.1519 | 4.2151 | 11.3093 | 3.7839 | 11.2012 | 3.9371
| Rel.Std.Dev. |  | 1.519% | 1.08% | 0.6366% | 1.71% | 0.5827% | 1.849% | 0.6107%
| PDM3 |  | 3.036% | -69.31% | 1.958% | 1.855% | -0.01724% | -6.742% | -0.7268%
| Z-Score (Absolute) |  | 1.04 | 23.65 | 0.67 | 0.64 | 0 | 2.3 | 0.24
| Analytical results for Be in OREAS 143 (Certified Value 3.14 ppm) | [13 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 13 | 14 | 17 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | BF*OES/MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | 3 | < 2 | 3 | 4 | 3.2096 | < 5 | 2.8441 | 4.9 |  |  |  |  | 1
| 1 | 2 | 3 | < 2 | 3 | 3 | 3.2815 | < 5 | 3.712 | 5.4 |  |  |  |  | #N/A
| 1 | 3 | 3 | < 2 | 3 | 3 | 3.2815 | < 5 | 3.8564 | 5.5 |  |  |  |  | 16
| 1 | 4 | 3 | < 2 | 3 | 3 | 3.0967 | < 5 | 3.2164 | 3.6 |  |  |  |  | 3.1441
| 1 | 5 | 3 | < 2 | 3 | 4 | 3.1035 | < 5 | 2.3922 | 5.6 |  |  |  |  | 34
| 1 | 6 | 3 | < 2 | 3 | 3 | 3.2917 | < 5 | 3.0389 | 5.7
| Mean |  | 3 |  | 3 | 3.3333 | 3.2107 |  | 3.1767 | 5.1167
| Median |  | 3 |  | 3 | 3 | 3.2455 |  | 3.1276 | 5.45
| Std Dev. |  | 0 |  | 0 | 0.5164 | 0.090637 |  | 0.54681 | 0.79352
| Rel.Std.Dev. |  | 0% |  | 0% | 15.49% | 2.823% |  | 17.21% | 15.51%
| PDM3 |  | -4.585% |  | -4.585% | 6.017% | 2.118% |  | 1.034% | 62.74%
| Z-Score (Absolute) |  | 0.32 | 7.67 | 0.32 | 0.9 | 0.45 | 2.16 | 0.32 | 7.44
| Analytical results for Bi in OREAS 143 (Indicative Value 23.7 ppm) | [13 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 04 | 05 | [8 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | [8 empty] | 1
1
| 1 | 1 | 24.1 | 15.8 | 23.8 | 23 | [8 empty] | 1
| 1 | 2 | 24.4 | 16.6 | 23.2 | 24 | [8 empty] | 1
| 1 | 3 | 24.6 | 15.7 | 22.9 | 23 | [8 empty] | 16
| 1 | 4 | 23.7 | 16.3 | 22.4 | 23 | [8 empty] | 23.65
| 1 | 5 | 24.8 | 16.6 | 24.4 | 23 | [8 empty] | 20
| 1 | 6 | 24.3 | 15.2 | 23.1 | 24
| Mean |  | 24.3167 | 16.0333 | 23.3 | 23.3333
| Median |  | 24.35 | 16.05 | 23.15 | 23
| Std Dev. |  | 0.38687 | 0.56095 | 0.70427 | 0.5164
| Rel.Std.Dev. |  | 1.591% | 3.499% | 3.023% | 2.213%
| PDM3 |  | 2.819% | -32.21% | -1.48% | -1.339%
| Z-Score (Absolute) |  | 1.33 | 9.66 | 0.02 | 0.02
| Analytical results for Ca in OREAS 143 (Certified Value 7.05 wt.%) | [13 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS |  |  | 2
3
| 1 | 1 | 7.2 | 6.1 | 7 | 6.93 | 7.397 | 7.111 | 7.111 | 7.2596 | 7.035 | 6.5494 |  |  | 1
| 1 | 2 | 7 | 6.18 | 6.9 | 7.22 | 7.033 | 6.911 | 7.147 | 7.2397 | 6.662 | 6.6267 |  |  | #N/A
| 1 | 3 | 7.2 | 6.2 | 7 | 7.52 | 7.54 | 6.861 | 6.797 | 7.2444 | 7.275 | 6.7524 |  |  | 16
| 1 | 4 | 7.4 | 6.11 | 7.2 | 7 | 7.068 | 7.047 | 6.947 | 7.2638 | 6.952 | 6.5672 |  |  | 7.0494
| 1 | 5 | 7.1 | 6.2 | 7 | 7.08 | 7.118 | 6.947 | 7.504 | 7.214 | 6.99 | 6.7325 |  |  | 35
| 1 | 6 | 7.1 | 6.19 | 7 | 7.02 | 7.29 | 6.539 | 7.254 | 7.2077 | 6.993 | 6.6104
| Mean |  | 7.1667 | 6.1633 | 7.0167 | 7.1283 | 7.241 | 6.9027 | 7.1267 | 7.2382 | 6.9845 | 6.6398
| Median |  | 7.15 | 6.185 | 7 | 7.05 | 7.204 | 6.929 | 7.129 | 7.242 | 6.9915 | 6.6185
| Std Dev. |  | 0.13663 | 0.045898 | 0.098319 | 0.21526 | 0.20222 | 0.20013 | 0.24504 | 0.023108 | 0.19601 | 0.084565
| Rel.Std.Dev. |  | 1.906% | 0.7447% | 1.401% | 3.02% | 2.793% | 2.899% | 3.438% | 0.3193% | 2.806% | 1.274%
| PDM3 |  | 1.663% | -12.57% | -0.4643% | 1.12% | 2.718% | -2.082% | 1.096% | 2.678% | -0.9206% | -5.811%
| Z-Score (Absolute) |  | 0.49 | 4.68 | 0.28 | 0.29 | 0.87 | 0.87 | 0.28 | 0.86 | 0.45 | 2.23
| Analytical results for Cd in OREAS 143 (Indicative Value 29.3 ppm) | [13 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 12 | 13 | [7 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | [7 empty] | 1
1
| 1 | 1 | < 1 | < 10 | 56 | 2.768 | < 10 | [7 empty] | 1
| 1 | 2 | < 1 | < 10 | 55 | 3.8 | < 10 | [7 empty] | 2
| 1 | 3 | < 1 | < 10 | 57 | 3.668 | < 10 | [7 empty] | 16
| 1 | 4 | < 1 | < 10 | 54 | 3.536 | < 10 | [7 empty] | 29.3003
| 1 | 5 | < 1 | < 10 | 56 | 3.644 | < 10 | [7 empty] | 21
| 1 | 6 | < 1 | < 10 | 53 | 3.188 | < 10
| Mean |  |  |  | 55.1667 | 3.434
| Median |  |  |  | 55.5 | 3.59
| Std Dev. |  |  |  | 1.472 | 0.3866
| Rel.Std.Dev. |  |  |  | 2.668% | 11.26%
| PDM3 |  |  |  | 88.28% | -88.28%
| Z-Score (Absolute) |  | 1.94 | 0 | 21.6 | 0.67 | 0
| Analytical results for Ce in OREAS 143 (Certified Value 67 ppm) | [13 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | BF*OES/MS | PF*MS | [6 empty] | 0
1
| 1 | 1 | 71.1 | 61.1 | 65.4 | 63.6 | 69.4324 | 63 | [6 empty] | 1
| 1 | 2 | 70.5 | 65.8 | 65.1 | 67.4 | 72.1465 | 63 | [6 empty] | 3
| 1 | 3 | 72.3 | 61.6 | 67.9 | 66.4 | 72.4383 | 64 | [6 empty] | 16
| 1 | 4 | 72.4 | 64 | 66.1 | 66.1 | 68.0197 | 62 | [6 empty] | 66.9466
| 1 | 5 | 71.5 | 62.6 | 66.9 | 68.3 | 73.6092 | 63 | [6 empty] | 36
| 1 | 6 | 68.7 | 62.5 | 66.4 | 67.9 | 73.8304 | 64
| Mean |  | 71.0833 | 62.9333 | 66.3 | 66.6167 | 71.5794 | 63.1667
| Median |  | 71.3 | 62.55 | 66.25 | 66.9 | 72.2924 | 63
| Std Dev. |  | 1.3717 | 1.7201 | 1.0218 | 1.7034 | 2.3464 | 0.75277
| Rel.Std.Dev. |  | 1.93% | 2.733% | 1.541% | 2.557% | 3.278% | 1.192%
| PDM3 |  | 6.179% | -5.995% | -0.9658% | -0.4928% | 6.92% | -5.646%
| Z-Score (Absolute) |  | 0.92 | 0.7 | 0.03 | 0.03 | 1.01 | 0.65
| Analytical results for Co in OREAS 143 (Certified Value 13.6 ppm) | [13 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 |  |  |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | BF*OES/MS |  |  |  | 1
2
| 1 | 1 | 14.5 | 13 | 12.9 | < 20 | 20 | < 20 | 11.6437 | 15 | 14.0336 |  |  |  | 1
| 1 | 2 | 14.3 | 14 | 14.4 | < 20 | < 20 | < 20 | 11.67 | 15 | 14.2873 |  |  |  | #N/A
| 1 | 3 | 14.5 | 13 | 12.1 | 20 | < 20 | 20 | 11.6212 | 15 | 14.1369 |  |  |  | 16
| 1 | 4 | 14.3 | 13 | 12.9 | < 20 | < 20 | < 20 | 11.875 | 15 | 14.4134 |  |  |  | 13.6125
| 1 | 5 | 14.3 | 13 | 14.1 | < 20 | < 20 | 30 | 11.85 | 13 | 14.3421 |  |  |  | 37
| 1 | 6 | 14.5 | 13 | 14.3 | < 20 | < 20 | < 20 | 11.2637 | 15 | 14.8117
| Mean |  | 14.4 | 13.1667 | 13.45 | 20 | 20 | 25 | 11.654 | 14.6667 | 14.3375
| Median |  | 14.4 | 13 | 13.5 | 20 | 20 | 25 | 11.6569 | 15 | 14.3147
| Std Dev. |  | 0.10954 | 0.40825 | 0.94604 |  |  | 7.0711 | 0.21954 | 0.8165 | 0.27048
| Rel.Std.Dev. |  | 0.7607% | 3.101% | 7.034% |  |  | 28.28% | 1.884% | 5.567% | 1.887%
| PDM3 |  | 5.785% | -3.275% | -1.193% | 46.92% | 46.92% | 83.66% | -14.39% | 7.744% | 5.326%
| Z-Score (Absolute) |  | 0.53 | 0.16 | 0 | 0.99 | 0.99 | 0.86 | 1 | 0.67 | 0.49
| Analytical results for Cr in OREAS 143 (Certified Value 74 ppm) | [13 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 13 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | [6 empty] | 0
0
| 1 | 1 | 70 | 90 | 70 | 140 | 70 | 79 | [6 empty] | 1
| 1 | 2 | 70 | 80 | 70 | 140 | 70 | 64 | [6 empty] | #N/A
| 1 | 3 | 70 | 90 | 70 | 210 | 70 | 71 | [6 empty] | 16
| 1 | 4 | 80 | 90 | 70 | 140 | 70 | 77 | [6 empty] | 74.3025
| 1 | 5 | 70 | 90 | 70 | 140 | 70 | 76 | [6 empty] | 38
| 1 | 6 | 80 | 90 | 70 | 140 | 70 | 71
| Mean |  | 73.3333 | 88.3333 | 70 | 151.6667 | 70 | 73
| Median |  | 70 | 90 | 70 | 140 | 70 | 73.5
| Std Dev. |  | 5.164 | 4.0825 | 0 | 28.5774 | 0 | 5.4772
| Rel.Std.Dev. |  | 7.042% | 4.622% | 0% | 18.84% | 0% | 7.503%
| PDM3 |  | -1.304% | 18.88% | -5.79% | 104.1% | -5.79% | -1.753%
| Z-Score (Absolute) |  | 0.02 | 2.16 | 0.67 | 10.67 | 0.67 | 0.02
| Analytical results for Cs in OREAS 143 (Certified Value 9.38 ppm) | [13 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 04 | 05 | 12 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 9.4 | 9.4 | 9.1 | 7.0617 | 10.0132 | 10 | [6 empty] | 1
| 1 | 2 | 9.3 | 9.32 | 9.6 | 6.9392 | 9.0849 | 9 | [6 empty] | 6
| 1 | 3 | 9.1 | 9.18 | 10.2 | 7.5709 | 9.3149 | 10 | [6 empty] | 16
| 1 | 4 | 9.3 | 9.24 | 9.5 | 6.907 | 9.503 | 8 | [6 empty] | 9.3816
| 1 | 5 | 9.9 | 9.18 | 9.3 | 7.0921 | 9.7339 | 17 | [6 empty] | 39
| 1 | 6 | 9 | 9.38 | 9.3 | 7.4368 | 9.5974 | 11
| Mean |  | 9.3333 | 9.2833 | 9.5 | 7.168 | 9.5412 | 10.8333
| Median |  | 9.3 | 9.28 | 9.4 | 7.0769 | 9.5502 | 10
| Std Dev. |  | 0.31411 | 0.097502 | 0.38471 | 0.2728 | 0.32372 | 3.1885
| Rel.Std.Dev. |  | 3.365% | 1.05% | 4.05% | 3.806% | 3.393% | 29.43%
| PDM3 |  | -0.5142% | -1.047% | 1.262% | -23.6% | 1.702% | 15.47%
| Z-Score (Absolute) |  | 0.44 | 0.7 | 0.44 | 11.76 | 0.65 | 7.41
| Analytical results for Cu in OREAS 143 (Certified Value 333 ppm) | [13 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 11 | 12 | 13 | 17 |  | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES |  | 0
0
| 1 | 1 | 334 | 309.5 | 340 | 337 | 330 | 340 | 330 | 330 | 350.7271 | 308 | 328 |  | 1
| 1 | 2 | 426 | 307.6 | 350 | 332 | 310 | 340 | 320 | 330 | 345.16 | 306 | 339 |  | #N/A
| 1 | 3 | 332 | 313.4 | 340 | 329 | 310 | 340 | 310 | 330 | 349.4127 | 307 | 339 |  | 16
| 1 | 4 | 335 | 311.7 | 320 | 340 | 310 | 330 | 320 | 340 | 347.0773 | 306 | 335 |  | 332.5379
| 1 | 5 | 340 | 312.6 | 340 | 336 | 320 | 340 | 330 | 330 | 351.6357 | 307 | 335 |  | 40
| 1 | 6 | 325 | 313.2 | 320 | 337 | 310 | 320 | 370 | 340 | 340.8326 | 308 | 344
| Mean |  | 348.6667 | 311.3333 | 335 | 335.1667 | 315 | 335 | 330 | 333.3333 | 347.4742 | 307 | 336.6667
| Median |  | 334.5 | 312.15 | 340 | 336.5 | 310 | 340 | 325 | 330 | 348.245 | 307 | 337
| Std Dev. |  | 38.1977 | 2.3149 | 12.2474 | 3.9707 | 8.3666 | 8.3666 | 20.9762 | 5.164 | 4.0319 | 0.89443 | 5.3914
| Rel.Std.Dev. |  | 10.96% | 0.7435% | 3.656% | 1.185% | 2.656% | 2.497% | 6.356% | 1.549% | 1.16% | 0.2913% | 1.601%
| PDM3 |  | 4.85% | -6.377% | 0.7404% | 0.7905% | -5.274% | 0.7404% | -0.7632% | 0.2392% | 4.492% | -7.68% | 1.242%
| Z-Score (Absolute) |  | 1.84 | 3.19 | 0 | 0.02 | 2.7 | 0 | 0.67 | 0.22 | 1.68 | 3.78 | 0.22
| Analytical results for Dy in OREAS 143 (Certified Value 5.01 ppm) | [13 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | BF*OES/MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 4.8 | 3.4 | 4.9 | 5 | 5.4533 | 5 | [6 empty] | 1
| 1 | 2 | 4.9 | 3.3 | 4.6 | 4.8 | 5.4611 | 5 | [6 empty] | #N/A
| 1 | 3 | 5.1 | 3.4 | 4.6 | 5 | 5.397 | 5 | [6 empty] | 16
| 1 | 4 | 5.3 | 3.3 | 4.7 | 5.2 | 5.3214 | 5 | [6 empty] | 5.0102
| 1 | 5 | 5.2 | 3.3 | 4.7 | 5 | 5.3812 | 5 | [6 empty] | 41
| 1 | 6 | 5.1 | 3.3 | 4.7 | 4.8 | 4.8921 | 5
| Mean |  | 5.0667 | 3.3333 | 4.7 | 4.9667 | 5.3177 | 5
| Median |  | 5.1 | 3.3 | 4.7 | 5 | 5.3891 | 5
| Std Dev. |  | 0.18619 | 0.05164 | 0.10954 | 0.15055 | 0.21468 | 0
| Rel.Std.Dev. |  | 3.675% | 1.549% | 2.331% | 3.031% | 4.037% | 0%
| PDM3 |  | 1.127% | -33.47% | -6.191% | -0.8689% | 6.137% | -0.2036%
| Z-Score (Absolute) |  | 0.31 | 6.07 | 1.04 | 0.06 | 1.23 | 0.06
| Analytical results for Er in OREAS 143 (Certified Value 2.75 ppm) | [13 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 2.8 | 1.7 | 2.7 | 2.5 | 2.1068 | 3.2194 | 3 | [5 empty] | 1
| 1 | 2 | 2.7 | 1.7 | 2.7 | 2.6 | 2.0172 | 3.0908 | 3 | [5 empty] | #N/A
| 1 | 3 | 3 | 1.7 | 2.7 | 3 | 2.0679 | 3.0706 | 3 | [5 empty] | 16
| 1 | 4 | 2.9 | 1.7 | 2.6 | 2.8 | 2.0834 | 3.002 | 3 | [5 empty] | 2.7537
| 1 | 5 | 3.2 | 1.6 | 2.7 | 2.8 | 2.0391 | 3.2429 | 3 | [5 empty] | 42
| 1 | 6 | 2.8 | 1.7 | 2.6 | 3 | 2.1814 | 2.9125 | 3
| Mean |  | 2.9 | 1.6833 | 2.6667 | 2.7833 | 2.0826 | 3.0897 | 3
| Median |  | 2.85 | 1.7 | 2.7 | 2.8 | 2.0756 | 3.0807 | 3
| Std Dev. |  | 0.17889 | 0.040825 | 0.05164 | 0.20412 | 0.05785 | 0.1263 | 0
| Rel.Std.Dev. |  | 6.168% | 2.425% | 1.936% | 7.334% | 2.778% | 4.088% | 0%
| PDM3 |  | 5.312% | -38.87% | -3.161% | 1.075% | -24.37% | 12.2% | 8.943%
| Z-Score (Absolute) |  | 0.36 | 3.42 | 0.36 | 0 | 2.18 | 0.95 | 0.67
| Analytical results for Eu in OREAS 143 (Certified Value 1.22 ppm) | [13 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 1.2 | 1.1 | 1.1 | 1.3 | 1.0758 | 1.4783 | < 2 | [5 empty] | 1
| 1 | 2 | 1.2 | 1.1 | 1.2 | 1.3 | 1.1249 | 1.344 | < 2 | [5 empty] | #N/A
| 1 | 3 | 1.3 | 1.1 | 1.1 | 1 | 1.2177 | 1.5312 | < 2 | [5 empty] | 16
| 1 | 4 | 1.2 | 1.1 | 1.2 | 1.3 | 1.0636 | 1.3356 | < 2 | [5 empty] | 1.2193
| 1 | 5 | 1.2 | 1.1 | 1.1 | 1.3 | 1.09 | 1.4056 | < 2 | [5 empty] | 43
| 1 | 6 | 1.2 | 1.1 | 1.2 | 1.4 | 1.1712 | 1.3371 | < 2
| Mean |  | 1.2167 | 1.1 | 1.15 | 1.2667 | 1.1239 | 1.4053
| Median |  | 1.2 | 1.1 | 1.15 | 1.3 | 1.1074 | 1.3748
| Std Dev. |  | 0.040825 | 2.4324e-16 | 0.054772 | 0.13663 | 0.060229 | 0.082998
| Rel.Std.Dev. |  | 3.355% | 2.211e-14% | 4.763% | 10.79% | 5.359% | 5.906%
| PDM3 |  | -0.2163% | -9.785% | -5.684% | 3.884% | -7.827% | 15.25%
| Z-Score (Absolute) |  | 0.67 | 0.51 | 0 | 1.18 | 0.26 | 2.58 | 1.52
| Analytical results for Fe in OREAS 143 (Certified Value 4.85 wt.%) | [13 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 11 | 12 | 13 | 14 | 17 | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*MS | 2
3
| 1 | 1 | 4.89 | 4.29 | 4.84 | 4.87 | 4.714 | 4.938 | 4.7 | 4.978 | 4.8593 | 4.845 | 4.8843 | 4.7064 | 1
| 1 | 2 | 4.84 | 4.25 | 4.83 | 4.99 | 4.686 | 4.91 | 4.819 | 4.912 | 4.8476 | 4.584 | 4.9659 | 4.8595 | #N/A
| 1 | 3 | 4.87 | 4.28 | 4.82 | 5.02 | 4.616 | 5.036 | 4.707 | 4.979 | 4.8636 | 4.965 | 4.9635 | 4.8313 | 16
| 1 | 4 | 5.04 | 4.35 | 4.85 | 4.81 | 4.679 | 4.889 | 4.749 | 5.026 | 4.8764 | 4.711 | 4.9799 | 4.7701 | 4.8545
| 1 | 5 | 4.87 | 4.32 | 4.85 | 4.98 | 4.756 | 4.959 | 4.938 | 4.89 | 4.8334 | 4.738 | 5.0102 | 4.776 | 44
| 1 | 6 | 4.78 | 4.32 | 4.8 | 4.89 | 4.686 | 4.721 | 4.952 | 4.979 | 4.8487 | 4.747 | 4.9909 | 4.8788
| Mean |  | 4.8817 | 4.3017 | 4.8317 | 4.9267 | 4.6895 | 4.9088 | 4.8108 | 4.9607 | 4.8548 | 4.765 | 4.9658 | 4.8037
| Median |  | 4.87 | 4.305 | 4.835 | 4.935 | 4.686 | 4.924 | 4.784 | 4.9785 | 4.854 | 4.7425 | 4.9729 | 4.8037
| Std Dev. |  | 0.086583 | 0.035449 | 0.019408 | 0.082138 | 0.045955 | 0.10509 | 0.1123 | 0.050206 | 0.014919 | 0.12894 | 0.043469 | 0.06462
| Rel.Std.Dev. |  | 1.774% | 0.8241% | 0.4017% | 1.667% | 0.98% | 2.141% | 2.334% | 1.012% | 0.3073% | 2.706% | 0.8754% | 1.345%
| PDM3 |  | 0.5592% | -11.39% | -0.4708% | 1.486% | -3.399% | 1.119% | -0.9% | 2.187% | 0.006389% | -1.844% | 2.292% | -1.047%
| Z-Score (Absolute) |  | 0.36 | 5.07 | 0.11 | 0.78 | 1.44 | 0.62 | 0.3 | 1.1 | 0.11 | 0.73 | 1.15 | 0.37
| Analytical results for Ga in OREAS 143 (Certified Value 17.9 ppm) | [13 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 18 | 18.5 | 18 | 19.7 | 16.882 | 17.9788 | [6 empty] | 1
| 1 | 2 | 16 | 18.6 | 18 | 19.6 | 16.644 | 16.8082 | [6 empty] | 7
| 1 | 3 | 19 | 18.7 | 18 | 19.3 | 16.698 | 15.6482 | [6 empty] | 16
| 1 | 4 | 18 | 18.4 | 18 | 20.5 | 16.48 | 16.5335 | [6 empty] | 17.8989
| 1 | 5 | 18 | 18.5 | 18 | 21.2 | 16.931 | 15.316 | [6 empty] | 45
| 1 | 6 | 18 | 18.8 | 18 | 19.7 | 16.581 | 16.8007
| Mean |  | 17.8333 | 18.5833 | 18 | 20 | 16.7027 | 16.5142
| Median |  | 18 | 18.55 | 18 | 19.7 | 16.671 | 16.6671
| Std Dev. |  | 0.98319 | 0.1472 | 0 | 0.70993 | 0.17444 | 0.94864
| Rel.Std.Dev. |  | 5.513% | 0.7921% | 0% | 3.55% | 1.044% | 5.744%
| PDM3 |  | -0.3665% | 3.824% | 0.5647% | 11.74% | -6.683% | -7.736%
| Z-Score (Absolute) |  | 0.06 | 0.48 | 0.06 | 1.49 | 0.87 | 1.01
| Analytical results for Gd in OREAS 143 (Certified Value 5.33 ppm) | [13 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 5.5 | 4.7 | 5.3 | 6.1 | 4.326 | 5.8353 | 6 | [5 empty] | 1
| 1 | 2 | 5.8 | 4.8 | 5.6 | 5.8 | 4.1877 | 5.7971 | 6 | [5 empty] | #N/A
| 1 | 3 | 6 | 4.7 | 5.6 | 5.8 | 4.4858 | 5.2855 | 5 | [5 empty] | 16
| 1 | 4 | 5.9 | 4.6 | 5.3 | 5.8 | 4.4835 | 5.7354 | 5 | [5 empty] | 5.3284
| 1 | 5 | 5.8 | 4.7 | 5.7 | 6.5 | 4.3956 | 5.3976 | 5 | [5 empty] | 46
| 1 | 6 | 5.4 | 4.7 | 5.3 | 6.2 | 4.62 | 5.6729 | 6
| Mean |  | 5.7333 | 4.7 | 5.4667 | 6.0333 | 4.4164 | 5.6206 | 5.5
| Median |  | 5.8 | 4.7 | 5.45 | 5.95 | 4.4396 | 5.7042 | 5.5
| Std Dev. |  | 0.23381 | 0.063246 | 0.18619 | 0.28752 | 0.14945 | 0.2259 | 0.54772
| Rel.Std.Dev. |  | 4.078% | 1.346% | 3.406% | 4.765% | 3.384% | 4.019% | 9.959%
| PDM3 |  | 7.6% | -11.79% | 2.595% | 13.23% | -17.12% | 5.484% | 3.221%
| Z-Score (Absolute) |  | 0.38 | 1.67 | 0.15 | 0.97 | 2.24 | 0.15 | NA
Results from laboratory 17 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 143 (Indicative Value 1.78 ppm) | [13 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 04 | 05 | 14 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | BF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 2 | 1.7 | 1.1735 | [9 empty] | 1
| 1 | 2 | 2 | 1.9 | 1.5101 | [9 empty] | 8
| 1 | 3 | 2 | 2 | 1.1871 | [9 empty] | 16
| 1 | 4 | 2 | 2 | 1.6868 | [9 empty] | 1.7816
| 1 | 5 | 2 | 2 | 1.6648 | [9 empty] | 22
| 1 | 6 | 2 | 1.7 | 1.5459
| Mean |  | 2 | 1.8833 | 1.4614
| Median |  | 2 | 1.95 | 1.528
| Std Dev. |  | 0 | 0.1472 | 0.22795
| Rel.Std.Dev. |  | 0% | 7.816% | 15.6%
| PDM3 |  | 12.26% | 5.712% | -17.97%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.44
| Analytical results for Hf in OREAS 143 (Indicative Value 4.93 ppm) | [13 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 05 | 14 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | BF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 6 | 3.5 | < 10 | 5.4578 | [8 empty] | 1
| 1 | 2 | 5.4 | 3.3 | < 10 | 5.4824 | [8 empty] | 17
| 1 | 3 | 5.8 | 3.4 | < 10 | 5.5443 | [8 empty] | 16
| 1 | 4 | 6.5 | 3.6 | < 10 | 5.2162 | [8 empty] | 4.9282
| 1 | 5 | 5.5 | 3.4 | < 10 | 5.4685 | [8 empty] | 23
| 1 | 6 | 6 | 3.5 | < 10 | 5.6386
| Mean |  | 5.8667 | 3.45 |  | 5.468
| Median |  | 5.9 | 3.45 |  | 5.4754
| Std Dev. |  | 0.39833 | 0.10488 |  | 0.14046
| Rel.Std.Dev. |  | 6.79% | 3.04% |  | 2.569%
| PDM3 |  | 19.04% | -29.99% |  | 10.95%
| Z-Score (Absolute) |  | 0.98 | 2.78 | 0.36 | 0.36
| Analytical results for Hg in OREAS 143 (Indicative Value < 5 ppm) | [13 empty] | Indicative
Hg | Round | Replicate | Lab | [11 empty] | 1
| No | No | 02 | [11 empty] | ppm
|  |  | PF*OES | [11 empty] | 2
2
| 1 | 1 | < 5 | [11 empty] | 1
| 1 | 2 | < 5 | [11 empty] | 18
| 1 | 3 | < 5 | [11 empty] | 16
| 1 | 4 | < 5 | [11 empty] | < 5
| 1 | 5 | < 5 | [11 empty] | 24
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 143 (Certified Value 0.96 ppm) | [13 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 1.1 | 0.6 | 1 | 1 | 0.7299 | 1.1596 | < 2 | [5 empty] | 1
| 1 | 2 | 1 | 0.6 | 1 | 1 | 0.7422 | 1.0063 | < 2 | [5 empty] | #N/A
| 1 | 3 | 1 | 0.6 | 1 | 0.9 | 0.7351 | 0.95932 | < 2 | [5 empty] | 16
| 1 | 4 | 1.1 | 0.6 | 1 | 1 | 0.7269 | 0.99499 | < 2 | [5 empty] | 0.95534
| 1 | 5 | 1.1 | 0.6 | 0.9 | 1 | 0.7229 | 1.0234 | < 2 | [5 empty] | 47
| 1 | 6 | 1.1 | 0.6 | 0.9 | 0.9 | 0.7379 | 1.1218 | < 2
| Mean |  | 1.0667 | 0.6 | 0.96667 | 0.96667 | 0.73248 | 1.0442
| Median |  | 1.1 | 0.6 | 1 | 1 | 0.7325 | 1.0148
| Std Dev. |  | 0.05164 | 0 | 0.05164 | 0.05164 | 0.0072151 | 0.078514
| Rel.Std.Dev. |  | 4.841% | 0% | 5.342% | 5.342% | 0.985% | 7.519%
| PDM3 |  | 11.65% | -37.2% | 1.185% | 1.185% | -23.33% | 9.304%
| Z-Score (Absolute) |  | 0.87 | 3.19 | 0 | 0 | 2.04 | 0.67 | 0.29
| Analytical results for In in OREAS 143 (Indicative Value 4.92 ppm) | [13 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 04 | 05 | 12 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 0.5 | 1.3 | 11.7 | 7.4088 | [8 empty] | 1
| 1 | 2 | 0.5 | 1.2 | 10.9 | 7.2454 | [8 empty] | 2
| 1 | 3 | 0.3 | 1.4 | 10.6 | 7.3682 | [8 empty] | 16
| 1 | 4 | 0.4 | 1.2 | 10.5 | 7.3998 | [8 empty] | 4.9225
| 1 | 5 | 0.3 | 1.4 | 10.5 | 7.228 | [8 empty] | 18
| 1 | 6 | 0.4 | 1.2 | 10.7 | 7.5494
| Mean |  | 0.4 | 1.2833 | 10.8167 | 7.3666
| Median |  | 0.4 | 1.25 | 10.65 | 7.384
| Std Dev. |  | 0.089443 | 0.098319 | 0.45789 | 0.11849
| Rel.Std.Dev. |  | 22.36% | 7.661% | 4.233% | 1.608%
| PDM3 |  | -91.87% | -73.93% | 119.7% | 49.65%
| Z-Score (Absolute) |  | 0.76 | 0.59 | 1.26 | 0.59
| Analytical results for K in OREAS 143 (Certified Value 2.38 wt.%) | [13 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 | 17 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*MS |  | 2
3
| 1 | 1 | 2.4 | 0.79 | 2.39 | 2.4 | 2.416 | 2.391 | 2.383 | 2.4593 | 2.41 | 2.2918 | 2.3675 |  | 1
| 1 | 2 | 2.4 | 0.79 | 2.38 | 2.4 | 2.374 | 2.441 | 2.424 | 2.4978 | 2.282 | 2.2834 | 2.3885 |  | #N/A
| 1 | 3 | 2.37 | 0.8 | 2.4 | 2.5 | 2.341 | 2.391 | 2.358 | 2.4897 | 2.476 | 2.2722 | 2.4109 |  | 16
| 1 | 4 | 2.43 | 0.8 | 2.41 | 2.3 | 2.358 | 2.333 | 2.358 | 2.4563 | 2.307 | 2.3144 | 2.3058 |  | 2.3788
| 1 | 5 | 2.37 | 0.8 | 2.37 | 2.4 | 2.391 | 2.341 | 2.449 | 2.4738 | 2.328 | 2.3006 | 2.3955 |  | 48
| 1 | 6 | 2.34 | 0.79 | 2.36 | 2.3 | 2.358 | 2.241 | 2.449 | 2.4981 | 2.347 | 2.2629 | 2.45
| Mean |  | 2.385 | 0.795 | 2.385 | 2.3833 | 2.373 | 2.3563 | 2.4035 | 2.4792 | 2.3583 | 2.2875 | 2.3864
| Median |  | 2.385 | 0.795 | 2.385 | 2.4 | 2.366 | 2.366 | 2.4035 | 2.4817 | 2.3375 | 2.2876 | 2.392
| Std Dev. |  | 0.031464 | 0.0054772 | 0.018708 | 0.075277 | 0.027011 | 0.06878 | 0.042712 | 0.018776 | 0.072146 | 0.018811 | 0.048137
| Rel.Std.Dev. |  | 1.319% | 0.689% | 0.7844% | 3.158% | 1.138% | 2.919% | 1.777% | 0.7574% | 3.059% | 0.8223% | 2.017%
| PDM3 |  | 0.2606% | -66.58% | 0.2606% | 0.1905% | -0.2439% | -0.9445% | 1.038% | 4.219% | -0.8605% | -3.836% | 0.318%
| Z-Score (Absolute) |  | 0.06 | 53.67 | 0.06 | 0 | 0.35 | 0.92 | 0.67 | 3.24 | 0.84 | 3.24 | 0.1
| Analytical results for La in OREAS 143 (Certified Value 32.5 ppm) | [13 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 13 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | [5 empty] | 1
2
| 1 | 1 | 34.8 | 30.3 | 33 | 35.4 | 30 | 50.7768 | 32 | [5 empty] | 1
| 1 | 2 | 34.8 | 30.8 | 32.5 | 35.1 | 28 | 48.0382 | 32 | [5 empty] | 3
| 1 | 3 | 36.4 | 30.5 | 33.6 | 35.5 | 29 | 48.1255 | 31 | [5 empty] | 16
| 1 | 4 | 35.9 | 30.5 | 32.8 | 35.2 | 29 | 46.0936 | 31 | [5 empty] | 32.5111
| 1 | 5 | 35.8 | 31.2 | 33 | 35.8 | 27 | 50.8946 | 33 | [5 empty] | 49
| 1 | 6 | 34.8 | 31.1 | 33.3 | 36.3 | 27 | 46.8137 | 33
| Mean |  | 35.4167 | 30.7333 | 33.0333 | 35.55 | 28.3333 | 48.457 | 32
| Median |  | 35.3 | 30.65 | 33 | 35.45 | 28.5 | 48.0818 | 32
| Std Dev. |  | 0.70545 | 0.36148 | 0.38297 | 0.44159 | 1.2111 | 1.9947 | 0.89443
| Rel.Std.Dev. |  | 1.992% | 1.176% | 1.159% | 1.242% | 4.274% | 4.116% | 2.795%
| PDM3 |  | 8.937% | -5.468% | 1.606% | 9.347% | -12.85% | 49.05% | -1.572%
| Z-Score (Absolute) |  | 0.67 | 0.65 | 0 | 0.71 | 1.33 | 4.36 | 0.29
| Analytical results for Li in OREAS 143 (Certified Value 59 ppm) | [13 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 04 | 05 | 12 | 13 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | [6 empty] | 0
1
| 1 | 1 | 57 | 58 | 62 | 57.7209 | 62 | 64 | [6 empty] | 1
| 1 | 2 | 59 | 57 | 64 | 57.0156 | 85 | 73 | [6 empty] | #N/A
| 1 | 3 | 59 | 58 | 65 | 57.5226 | 68 | 66 | [6 empty] | 16
| 1 | 4 | 56 | 57 | 63 | 57.5365 | 78 | 55 | [6 empty] | 59.0658
| 1 | 5 | 60 | 57 | 64 | 57.5515 | 84 | 53 | [6 empty] | 50
| 1 | 6 | 56 | 58 | 64 | 57.8278 | 70 | 56
| Mean |  | 57.8333 | 57.5 | 63.6667 | 57.5291 | 74.5 | 61.1667
| Median |  | 58 | 57.5 | 64 | 57.544 | 74 | 60
| Std Dev. |  | 1.7224 | 0.54772 | 1.0328 | 0.27942 | 9.2898 | 7.7825
| Rel.Std.Dev. |  | 2.978% | 0.9526% | 1.622% | 0.4857% | 12.47% | 12.72%
| PDM3 |  | -2.087% | -2.651% | 7.789% | -2.602% | 26.13% | 3.557%
| Z-Score (Absolute) |  | 0.57 | 0.68 | 1.42 | 0.67 | 5.09 | 0.57
| Analytical results for Lu in OREAS 143 (Indicative Value 0.38 ppm) | [13 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 12 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 0.4 | < 0.5 | 0.38 | 0.2914 | 0.41462 | < 2 | [6 empty] | 1
| 1 | 2 | 0.4 | < 0.5 | 0.39 | 0.2832 | 0.43898 | < 2 | [6 empty] | 13
| 1 | 3 | 0.4 | < 0.5 | 0.36 | 0.2932 | 0.42791 | < 2 | [6 empty] | 16
| 1 | 4 | 0.4 | < 0.5 | 0.37 | 0.265 | 0.42618 | < 2 | [6 empty] | 0.38064
| 1 | 5 | 0.5 | < 0.5 | 0.4 | 0.2821 | 0.44556 | < 2 | [6 empty] | 19
| 1 | 6 | 0.5 | < 0.5 | 0.35 | 0.3066 | 0.41057 | < 2
| Mean |  | 0.43333 |  | 0.375 | 0.28692 | 0.4273
| Median |  | 0.4 |  | 0.375 | 0.2873 | 0.42704
| Std Dev. |  | 0.05164 |  | 0.018708 | 0.013888 | 0.013506
| Rel.Std.Dev. |  | 11.92% |  | 4.989% | 4.84% | 3.161%
| PDM3 |  | 13.84% |  | -1.481% | -24.62% | 12.26%
| Z-Score (Absolute) |  | 0.3 | 1.39 | 0.24 | 1.05 | 0.24 | 5.52
| Analytical results for Mg in OREAS 143 (Certified Value 1.2 wt.%) | [13 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS |  |  | 2
3
| 1 | 1 | 1.23 | 0.9 | 1.21 | 1.21 | 1.194 | 1.218 | 1.2 | 1.1936 | 1.193 | 1.2232 |  |  | 1
| 1 | 2 | 1.24 | 0.91 | 1.2 | 1.26 | 1.17 | 1.236 | 1.2 | 1.1352 | 1.137 | 1.2214 |  |  | #N/A
| 1 | 3 | 1.23 | 0.91 | 1.21 | 1.25 | 1.17 | 1.2 | 1.182 | 1.1248 | 1.233 | 1.209 |  |  | 16
| 1 | 4 | 1.25 | 0.92 | 1.2 | 1.21 | 1.164 | 1.194 | 1.182 | 1.1937 | 1.175 | 1.22 |  |  | 1.1994
| 1 | 5 | 1.23 | 0.92 | 1.2 | 1.25 | 1.188 | 1.206 | 1.242 | 1.1 | 1.181 | 1.2242 |  |  | 51
| 1 | 6 | 1.21 | 0.91 | 1.19 | 1.23 | 1.182 | 1.152 | 1.23 | 1.106 | 1.165 | 1.2095
| Mean |  | 1.2317 | 0.91167 | 1.2017 | 1.235 | 1.178 | 1.201 | 1.206 | 1.1422 | 1.1807 | 1.2179
| Median |  | 1.23 | 0.91 | 1.2 | 1.24 | 1.176 | 1.203 | 1.2 | 1.13 | 1.178 | 1.2207
| Std Dev. |  | 0.013292 | 0.0075277 | 0.0075277 | 0.021679 | 0.011798 | 0.028249 | 0.024884 | 0.041797 | 0.03186 | 0.0068647
| Rel.Std.Dev. |  | 1.079% | 0.8257% | 0.6264% | 1.755% | 1.002% | 2.352% | 2.063% | 3.659% | 2.698% | 0.5636%
| PDM3 |  | 2.693% | -23.99% | 0.1913% | 2.971% | -1.782% | 0.1357% | 0.5526% | -4.764% | -1.56% | 1.545%
| Z-Score (Absolute) |  | 0.93 | 8.86 | 0.01 | 1.03 | 0.72 | 0.01 | 0.14 | 1.81 | 0.63 | 0.51
| Analytical results for Mn in OREAS 143 (Certified Value 0.046 wt.%) | [13 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 |  |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.04165 | 0.046 | 0.0457 | 0.046 | 0.046 | 0.046 | 0.045229 | 0.0452 | 0.047777 |  |  |  | 1
| 1 | 2 | 0.04154 | 0.047 | 0.0477 | 0.046 | 0.046 | 0.046 | 0.044959 | 0.0426 | 0.048683 |  |  |  | #N/A
| 1 | 3 | 0.04144 | 0.046 | 0.0465 | 0.046 | 0.054 | 0.046 | 0.045183 | 0.0452 | 0.04708 |  |  |  | 16
| 1 | 4 | 0.0419 | 0.045 | 0.0462 | 0.046 | 0.046 | 0.046 | 0.045325 | 0.0436 | 0.047966 |  |  |  | 0.04561
| 1 | 5 | 0.04179 | 0.046 | 0.0469 | 0.046 | 0.046 | 0.046 | 0.045217 | 0.0428 | 0.048526 |  |  |  | 52
| 1 | 6 | 0.04192 | 0.045 | 0.0462 | 0.046 | 0.046 | 0.046 | 0.045043 | 0.0442 | 0.047479
| Mean |  | 0.041707 | 0.045833 | 0.046533 | 0.046 | 0.047333 | 0.046 | 0.045159 | 0.043933 | 0.047918
| Median |  | 0.04172 | 0.046 | 0.04635 | 0.046 | 0.046 | 0.046 | 0.0452 | 0.0439 | 0.047871
| Std Dev. |  | 0.00019593 | 0.00075277 | 0.00069474 | 7.6012e-18 | 0.003266 | 7.6012e-18 | 0.00013377 | 0.0011361 | 0.00061185
| Rel.Std.Dev. |  | 0.4698% | 1.642% | 1.493% | 1.652e-14% | 6.9% | 1.652e-14% | 0.2962% | 2.586% | 1.277%
| PDM3 |  | -8.557% | 0.4905% | 2.025% | 0.8559% | 3.779% | 0.8559% | -0.9875% | -3.675% | 5.062%
| Z-Score (Absolute) |  | 2.49 | 0.33 | 0.03 | 0 | 0.67 | 0 | 0.68 | 1.32 | 0.76
| Analytical results for Mo in OREAS 143 (Indicative Value 2.94 ppm) | [13 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 13 | [8 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | [8 empty] | 2
2
| 1 | 1 | < 1 | 1 | 3 | 11 | [8 empty] | 1
| 1 | 2 | < 1 | 1 | 3 | < 10 | [8 empty] | 14
| 1 | 3 | < 1 | 1 | 3 | < 10 | [8 empty] | 16
| 1 | 4 | < 1 | 1 | 3 | < 10 | [8 empty] | 2.9444
| 1 | 5 | < 1 | 1 | 2 | < 10 | [8 empty] | 20
| 1 | 6 | < 1 | 1 | 3 | < 10
| Mean |  |  | 1 | 2.8333 | 11
| Median |  |  | 1 | 3 | 11
| Std Dev. |  |  | 0 | 0.40825
| Rel.Std.Dev. |  |  | 0% | 14.41%
| PDM3 |  |  | -66.04% | -3.774% | 273.6%
| Z-Score (Absolute) |  | 0.82 | 0.53 | 0.53 | 2.36
| Analytical results for Na in OREAS 143 (Indicative Value 0.512 wt.%) | [13 empty] | Indicative
Na | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 02 | 14 | [10 empty] | wt.%
|  |  | PF*OES | BF*OES/MS | [10 empty] | 3
3
| 1 | 1 | 0.08 | 0.91832 | [10 empty] | 1
| 1 | 2 | 0.08 | 0.92867 | [10 empty] | 15
| 1 | 3 | 0.09 | 0.94464 | [10 empty] | 16
| 1 | 4 | 0.08 | 0.94079 | [10 empty] | 0.51157
| 1 | 5 | 0.09 | 0.9634 | [10 empty] | 21
| 1 | 6 | 0.09 | 0.93307
| Mean |  | 0.085 | 0.93815
| Median |  | 0.085 | 0.93693
| Std Dev. |  | 0.0054772 | 0.015462
| Rel.Std.Dev. |  | 6.444% | 1.648%
| PDM3 |  | -83.38% | 83.38%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 143 (Indicative Value 12 ppm) | [13 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 05 | 14 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | [9 empty] | 1
1
| 1 | 1 | 11.9 | 11.7 | 12.5173 | [9 empty] | 1
| 1 | 2 | 12 | 12.2 | 11.8417 | [9 empty] | 4
| 1 | 3 | 12 | 11.8 | 11.8339 | [9 empty] | 16
| 1 | 4 | 12 | 12.2 | 11.5721 | [9 empty] | 12.0228
| 1 | 5 | 11.9 | 12.3 | 11.8203 | [9 empty] | 22
| 1 | 6 | 12.2 | 12.4 | 12.2249
| Mean |  | 12 | 12.1 | 11.9684
| Median |  | 12 | 12.2 | 11.8378
| Std Dev. |  | 0.10954 | 0.28284 | 0.34066
| Rel.Std.Dev. |  | 0.9129% | 2.338% | 2.846%
| PDM3 |  | -0.1896% | 0.6422% | -0.4526%
| Z-Score (Absolute) |  | 0 | 2.13 | 0.67
| Analytical results for Nd in OREAS 143 (Certified Value 30.5 ppm) | [13 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 1
2
| 1 | 1 | 31.1 | 28.19 | 31 | 30.6 | 30.423 | 32.6851 | 28 | [5 empty] | 1
| 1 | 2 | 31.2 | 28.73 | 30.5 | 31 | 30.082 | 32.2553 | 29 | [5 empty] | #N/A
| 1 | 3 | 32 | 27.65 | 30 | 32.2 | 30.604 | 34.0217 | 28 | [5 empty] | 16
| 1 | 4 | 32 | 27.37 | 30.2 | 30.7 | 29.9875 | 33.1909 | 28 | [5 empty] | 30.4628
| 1 | 5 | 32 | 28.79 | 31.2 | 31.2 | 30.6035 | 31.8308 | 28 | [5 empty] | 53
| 1 | 6 | 31.1 | 29.01 | 30.3 | 32.7 | 30.659 | 32.3536 | 29
| Mean |  | 31.5667 | 28.29 | 30.5333 | 31.4 | 30.3932 | 32.7229 | 28.3333
| Median |  | 31.6 | 28.46 | 30.4 | 31.1 | 30.5132 | 32.5194 | 28
| Std Dev. |  | 0.4761 | 0.66753 | 0.47188 | 0.85557 | 0.29039 | 0.78206 | 0.5164
| Rel.Std.Dev. |  | 1.508% | 2.36% | 1.545% | 2.725% | 0.9555% | 2.39% | 1.823%
| PDM3 |  | 3.624% | -7.133% | 0.2316% | 3.077% | -0.2285% | 7.419% | -6.99%
| Z-Score (Absolute) |  | 0.67 | 1.46 | 0 | 0.57 | 0.09 | 1.43 | 1.44
| Analytical results for Ni in OREAS 143 (Certified Value 36.4 ppm) | [13 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 11 | 13 |  |  |  |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 1
1
| 1 | 1 | 27.26 | 40 | 40 | 30 | 70 | 30 | < 100 | 40 |  |  |  |  | 1
| 1 | 2 | 27.54 | 40 | 40 | 40 | 70 | 40 | < 100 | 32 |  |  |  |  | #N/A
| 1 | 3 | 27.43 | 40 | 40 | 40 | 800 | 40 | < 100 | 34 |  |  |  |  | 16
| 1 | 4 | 27.91 | 40 | 50 | 30 | 70 | 40 | < 100 | 33 |  |  |  |  | 36.4081
| 1 | 5 | 27.56 | 40 | 40 | 50 | 60 | 40 | < 100 | 31 |  |  |  |  | 54
| 1 | 6 | 27.99 | 40 | 40 | 30 | 50 | 40 | < 100 | 35
| Mean |  | 27.615 | 40 | 41.6667 | 36.6667 | 186.6667 | 38.3333 |  | 34.1667
| Median |  | 27.55 | 40 | 40 | 35 | 70 | 40 |  | 33.5
| Std Dev. |  | 0.28162 | 0 | 4.0825 | 8.165 | 300.5772 | 4.0825 |  | 3.1885
| Rel.Std.Dev. |  | 1.02% | 0% | 9.798% | 22.27% | 161% | 10.65% |  | 9.332%
| PDM3 |  | -24.15% | 9.866% | 14.44% | 0.7103% | 412.7% | 5.288% |  | -6.156%
| Z-Score (Absolute) |  | 2.08 | 0.15 | 0.45 | 0.45 | 26.52 | 0.15 | 1.95 | 0.9
| Analytical results for P in OREAS 143 (Certified Value 0.057 wt.%) | [13 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 12 | 13 | 14 | [6 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.06 | 0.05 | 0.06 | 0.054846 | 0.06 | 0.057824 | [6 empty] | 1
| 1 | 2 | 0.06 | 0.05 | 0.06 | 0.055177 | 0.054 | 0.057499 | [6 empty] | #N/A
| 1 | 3 | 0.05 | 0.05 | 0.06 | 0.054927 | 0.06 | 0.055725 | [6 empty] | 16
| 1 | 4 | 0.06 | 0.05 | 0.06 | 0.054917 | 0.054 | 0.058485 | [6 empty] | 0.05748
| 1 | 5 | 0.06 | 0.05 | 0.06 | 0.055127 | 0.058 | 0.055175 | [6 empty] | 55
| 1 | 6 | 0.06 | 0.05 | 0.06 | 0.05548 | 0.057 | 0.056232
| Mean |  | 0.058333 | 0.05 | 0.06 | 0.055079 | 0.057167 | 0.056823
| Median |  | 0.06 | 0.05 | 0.06 | 0.055027 | 0.0575 | 0.056866
| Std Dev. |  | 0.0040825 | 7.6012e-18 | 0 | 0.00023506 | 0.0027142 | 0.0013031
| Rel.Std.Dev. |  | 6.999% | 1.52e-14% | 0% | 0.4268% | 4.748% | 2.293%
| PDM3 |  | 1.484% | -13.01% | 4.383% | -4.178% | -0.5459% | -1.143%
| Z-Score (Absolute) |  | 0.55 | 2.9 | 1.25 | 0.79 | 0.07 | 0.07
| Analytical results for Pb in OREAS 143 (Indicative Value 40.3 ppm) | [13 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 11 | 13 |  |  |  | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  | 1
1
| 1 | 1 | 35 | 23.1 | 40 | 32.4 | < 100 | < 100 | 2200 | < 100 | 73 |  |  |  | 1
| 1 | 2 | 25 | 24.4 | 40 | 28.6 | 100 | < 100 | < 100 | < 100 | 49 |  |  |  | 6
| 1 | 3 | 29 | 24.2 | 40 | 29.2 | < 100 | < 100 | < 100 | < 100 | 61 |  |  |  | 16
| 1 | 4 | 30 | 25.2 | 30 | 29.2 | < 100 | 100 | < 100 | < 100 | 66 |  |  |  | 40.2667
| 1 | 5 | 33 | 24.7 | 30 | 32.7 | < 100 | < 100 | 100 | < 100 | 72 |  |  |  | 23
| 1 | 6 | 27 | 23.7 | 30 | 31.8 | < 100 | < 100 | 8400 | < 100 | 52
| Mean |  | 29.8333 | 24.2167 | 35 | 30.65 | 100 | 100 | 3566.6667 |  | 62.1667
| Median |  | 29.5 | 24.3 | 35 | 30.5 | 100 | 100 | 2200 |  | 63.5
| Std Dev. |  | 3.7103 | 0.7414 | 5.4772 | 1.8436 |  |  | 4315.476 |  | 10.0681
| Rel.Std.Dev. |  | 12.44% | 3.062% | 15.65% | 6.015% |  |  | 121% |  | 16.2%
| PDM3 |  | -25.91% | -39.86% | -13.08% | -23.88% | 148.3% | 148.3% | 8758% |  | 54.39%
| Z-Score (Absolute) |  | 0.91 | 1.16 | 0.67 | 0.87 | 0.37 | 0.37 | 79.04 | 0 | 0.55
| Analytical results for Pr in OREAS 143 (Certified Value 7.82 ppm) | [13 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | BF*OES/MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 8.5 | 6.5 | 7.66 | 7.9 | 8.7041 | 8 | [6 empty] | 1
| 1 | 2 | 8.4 | 6.5 | 7.69 | 8.2 | 8.7057 | 7 | [6 empty] | #N/A
| 1 | 3 | 8.7 | 6.3 | 7.93 | 7.9 | 8.7073 | 8 | [6 empty] | 16
| 1 | 4 | 8.7 | 6.2 | 7.64 | 8.1 | 8.582 | 7 | [6 empty] | 7.8221
| 1 | 5 | 8.5 | 6.6 | 7.78 | 8.4 | 8.5402 | 7 | [6 empty] | 56
| 1 | 6 | 8.4 | 6.6 | 7.72 | 8.2 | 8.3378 | 8
| Mean |  | 8.5333 | 6.45 | 7.7367 | 8.1167 | 8.5962 | 7.5
| Median |  | 8.5 | 6.5 | 7.705 | 8.15 | 8.643 | 7.5
| Std Dev. |  | 0.13663 | 0.16432 | 0.10671 | 0.19408 | 0.14565 | 0.54772
| Rel.Std.Dev. |  | 1.601% | 2.548% | 1.379% | 2.391% | 1.694% | 7.303%
| PDM3 |  | 9.092% | -17.54% | -1.093% | 3.765% | 9.896% | -4.118%
| Z-Score (Absolute) |  | 0.79 | 1.93 | 0.25 | 0.25 | 0.87 | 0.56
| Analytical results for Rb in OREAS 143 (Certified Value 151 ppm) | [13 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 04 | 05 | 12 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [6 empty] | 0
0
| 1 | 1 | 148.9 | 149 | 151 | 154.566 | 157.5051 | 159 | [6 empty] | 1
| 1 | 2 | 149.1 | 151 | 153 | 157.492 | 158.2179 | 144 | [6 empty] | 7
| 1 | 3 | 149.9 | 149 | 152 | 153.278 | 158.305 | 148 | [6 empty] | 16
| 1 | 4 | 149.7 | 151 | 151 | 152.464 | 156.4206 | 140 | [6 empty] | 151.2258
| 1 | 5 | 153.7 | 151 | 155 | 153.979 | 159.2292 | 162 | [6 empty] | 57
| 1 | 6 | 146.6 | 150 | 150 | 157.095 | 157.4885 | 144
| Mean |  | 149.65 | 150.1667 | 152 | 154.8123 | 157.861 | 149.5
| Median |  | 149.4 | 150.5 | 151.5 | 154.2725 | 157.8615 | 146
| Std Dev. |  | 2.308 | 0.98319 | 1.7889 | 2.0501 | 0.95255 | 8.9387
| Rel.Std.Dev. |  | 1.542% | 0.6547% | 1.177% | 1.324% | 0.6034% | 5.979%
| PDM3 |  | -1.042% | -0.7004% | 0.5119% | 2.372% | 4.388% | -1.141%
| Z-Score (Absolute) |  | 0.64 | 0.41 | 0.41 | 1.67 | 3.03 | 0.71
| Analytical results for Re in OREAS 143 (Indicative Value < 0.1 ppm) | [13 empty] | Indicative
Re | Round | Replicate | Lab | [11 empty] | 1
| No | No | 04 | [11 empty] | ppm
|  |  | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | < 0.1 | [11 empty] | 1
| 1 | 2 | < 0.1 | [11 empty] | 18
| 1 | 3 | < 0.1 | [11 empty] | 16
| 1 | 4 | < 0.1 | [11 empty] | < 0.1
| 1 | 5 | < 0.1 | [11 empty] | 24
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 143 (Certified Value 1.25 wt.%) | [13 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 17 |  |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 1.23 | 1.23 | 1.25 | 1.18 | 1.27 | 1.27 | 1.25 | 1.2228 | 1.2456 |  |  |  | 1
| 1 | 2 | 1.22 | 1.24 | 1.26 | 1.22 | 1.24 | 1.29 | 1.25 | 1.2139 | 1.2636 |  |  |  | 24
| 1 | 3 | 1.22 | 1.23 | 1.26 | 1.22 | 1.25 | 1.27 | 1.22 | 1.2523 | 1.2785 |  |  |  | 16
| 1 | 4 | 1.29 | 1.23 | 1.25 | 1.16 | 1.24 | 1.25 | 1.26 | 1.203 | 1.2242 |  |  |  | 1.2478
| 1 | 5 | 1.22 | 1.23 | 1.25 | 1.24 | 1.25 | 1.29 | 1.28 | 1.2144 | 1.2401 |  |  |  | 58
| 1 | 6 | 1.24 | 1.22 | 1.26 | 1.18 | 1.26 | 1.2 | 1.28 | 1.2092 | 1.2698
| Mean |  | 1.2367 | 1.23 | 1.255 | 1.2 | 1.2517 | 1.2617 | 1.2567 | 1.2193 | 1.2536
| Median |  | 1.225 | 1.23 | 1.255 | 1.2 | 1.25 | 1.27 | 1.255 | 1.2141 | 1.2546
| Std Dev. |  | 0.027325 | 0.0063246 | 0.0054772 | 0.030984 | 0.01169 | 0.033714 | 0.022509 | 0.01744 | 0.02046
| Rel.Std.Dev. |  | 2.21% | 0.5142% | 0.4364% | 2.582% | 0.934% | 2.672% | 1.791% | 1.43% | 1.632%
| PDM3 |  | -0.8926% | -1.427% | 0.5766% | -3.831% | 0.3095% | 1.111% | 0.7102% | -2.287% | 0.4671%
| Z-Score (Absolute) |  | 1.01 | 1.46 | 0.22 | 3.48 | 0 | 0.67 | 0.34 | 2.18 | 0.13
| Analytical results for Sb in OREAS 143 (Indicative Value 2.48 ppm) | [13 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 11 | 13 | 14 | [5 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES | BF*OES/MS | [5 empty] | 2
2
| 1 | 1 | 2.4 | < 5 | 2 | 3 | < 100 | < 50 | 2.8389 | [5 empty] | 1
| 1 | 2 | 2.2 | < 5 | 2.2 | 2 | < 100 | < 50 | 3.0862 | [5 empty] | 8
| 1 | 3 | 2.4 | < 5 | 2.1 | 3 | < 100 | < 50 | 2.9834 | [5 empty] | 16
| 1 | 4 | 2.4 | < 5 | 1.9 | < 2 | < 100 | < 50 | 3.0058 | [5 empty] | 2.4779
| 1 | 5 | 2.8 | < 5 | 2 | 3 | < 100 | < 50 | 3.2075 | [5 empty] | 18
| 1 | 6 | 1.9 | < 5 | 1.9 | 3 | < 100 | < 50 | 3.1489
| Mean |  | 2.35 |  | 2.0167 | 2.8 |  |  | 3.0451
| Median |  | 2.4 |  | 2 | 3 |  |  | 3.046
| Std Dev. |  | 0.29496 |  | 0.1169 | 0.44721 |  |  | 0.13167
| Rel.Std.Dev. |  | 12.55% |  | 5.797% | 15.97% |  |  | 4.324%
| PDM3 |  | -5.163% |  | -18.62% | 13% |  |  | 22.89%
| Z-Score (Absolute) |  | 0.21 | 0 | 0.67 | 0 | 66.27 | 31.39 | 0.76
| Analytical results for Sc in OREAS 143 (Indicative Value 9.82 ppm) | [13 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 13 | 14 | 17 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 10.4 | < 20 | 8 | 12.0649 | 11 | [7 empty] | 1
| 1 | 2 | 10.5 | < 20 | 6 | 11.7724 | 12 | [7 empty] | 1
| 1 | 3 | 10.2 | < 20 | 6 | 12.5253 | 12 | [7 empty] | 16
| 1 | 4 | 10.1 | < 20 | 7 | 11.6725 | 11 | [7 empty] | 9.8187
| 1 | 5 | 10.1 | < 20 | < 5 | 12.0367 | 11 | [7 empty] | 19
| 1 | 6 | 10.6 | < 20 | < 5 | 12.6765 | 12
| Mean |  | 10.3167 |  | 6.75 | 12.1247 | 11.5
| Median |  | 10.3 |  | 6.5 | 12.0508 | 11.5
| Std Dev. |  | 0.2137 |  | 0.95743 | 0.40123 | 0.54772
| Rel.Std.Dev. |  | 2.071% |  | 14.18% | 3.309% | 4.763%
| PDM3 |  | 5.072% |  | -31.25% | 23.49% | 17.12%
| Z-Score (Absolute) |  | 0 | 0.18 | 2.84 | 1.03 | 0.67
| Analytical results for Se in OREAS 143 (Indicative Value < 8 ppm) | [13 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 04 | 05 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 20 | < 8 | [10 empty] | 1
| 1 | 2 | < 20 | < 8 | [10 empty] | 22
| 1 | 3 | < 20 | < 8 | [10 empty] | 16
| 1 | 4 | < 20 | < 8 | [10 empty] | < 8
| 1 | 5 | < 20 | < 8 | [10 empty] | 20
| 1 | 6 | < 20 | < 8
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Si in OREAS 143 (Certified Value 25.8 wt.%) | [13 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 04 | 05 | 06 | 08 | 09 | 12 | 13 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 2
3
| 1 | 1 | 25.5 | 26.3 | 24.9 | 24.59 | 27.02 | 25.9 | 26.5616 | 25.309 |  |  |  |  | 1
| 1 | 2 | 25.6 | 26.4 | 26.5 | 24.12 | 27.72 | 26.08 | 26.9782 | 23.68 |  |  |  |  | #N/A
| 1 | 3 | 25.3 | 26.6 | 26.3 | 23.7 | 27.02 | 25.99 | 26.643 | 25.983 |  |  |  |  | 16
| 1 | 4 | 25.8 | 26.3 | 25.4 | 23.98 | 26.97 | 25.71 | 26.0207 | 24.582 |  |  |  |  | 25.8014
| 1 | 5 | 25.4 | 28.7 | 26.4 | 24.4 | 26.88 | 26.88 | 26.2823 | 24.955 |  |  |  |  | 59
| 1 | 6 | 24.8 | 26.5 | 25.4 | 24.12 | 25.94 | 26.51 | 26.8286 | 25.287
| Mean |  | 25.4 | 26.8 | 25.8167 | 24.1517 | 26.925 | 26.1783 | 26.5524 | 24.966
| Median |  | 25.45 | 26.45 | 25.85 | 24.12 | 26.995 | 26.035 | 26.6023 | 25.121
| Std Dev. |  | 0.34059 | 0.93808 | 0.66758 | 0.31282 | 0.56997 | 0.43458 | 0.35259 | 0.78184
| Rel.Std.Dev. |  | 1.341% | 3.5% | 2.586% | 1.295% | 2.117% | 1.66% | 1.328% | 3.132%
| PDM3 |  | -1.556% | 3.87% | 0.05909% | -6.394% | 4.355% | 1.461% | 2.911% | -3.238%
| Z-Score (Absolute) |  | 0.58 | 0.77 | 0.17 | 1.78 | 0.89 | 0.17 | 0.53 | 0.99
| Analytical results for Sm in OREAS 143 (Certified Value 6.13 ppm) | [13 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 12 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 6.5 | 5.4 | 6.2 | 4.4275 | 6.7222 | 6 | [6 empty] | 1
| 1 | 2 | 6.6 | 5.5 | 6.7 | 4.4014 | 6.1547 | 6 | [6 empty] | #N/A
| 1 | 3 | 6.6 | 5.3 | 6.4 | 4.81 | 6.1292 | 6 | [6 empty] | 16
| 1 | 4 | 6.6 | 5.3 | 6.6 | 4.5604 | 6.3245 | 6 | [6 empty] | 6.1289
| 1 | 5 | 6.4 | 5.4 | 6.6 | 4.5242 | 6.3031 | 6 | [6 empty] | 60
| 1 | 6 | 6.3 | 5.5 | 6.3 | 4.9147 | 6.0321 | 6
| Mean |  | 6.5 | 5.4 | 6.4667 | 4.6064 | 6.2776 | 6
| Median |  | 6.55 | 5.4 | 6.5 | 4.5423 | 6.2289 | 6
| Std Dev. |  | 0.12649 | 0.089443 | 0.19664 | 0.20948 | 0.24409 | 0
| Rel.Std.Dev. |  | 1.946% | 1.656% | 3.041% | 4.548% | 3.888% | 0%
| PDM3 |  | 6.056% | -11.89% | 5.512% | -24.84% | 2.427% | -2.102%
| Z-Score (Absolute) |  | 0.71 | 1.45 | 0.64 | 3 | 0.27 | 0.27
| Analytical results for Sn in OREAS 143 (Certified Value 0.238 wt.%) | [13 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 | 17 |  | wt.%
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*MS |  | 3
3
| 1 | 1 | 0.2601 | 0.20886 | 0.24 | 0.254 | 0.24 | 0.24 | 0.24 | 0.20983 | 0.2266 | 0.23263 | 0.2584 |  | 1
| 1 | 2 | 0.2566 | 0.21088 | 0.24 | 0.251 | 0.24 | 0.24 | 0.25 | 0.20716 | 0.2241 | 0.23376 | 0.2681 |  | 20
| 1 | 3 | 0.2567 | 0.21113 | 0.24 | 0.251 | 0.24 | 0.24 | 0.24 | 0.20243 | 0.225 | 0.23307 | 0.265 |  | 16
| 1 | 4 | 0.26 | 0.20979 | 0.24 | 0.248 | 0.24 | 0.24 | 0.25 | 0.20951 | 0.22 | 0.22777 | 0.2613 |  | 0.23755
| 1 | 5 | 0.2537 | 0.21006 | 0.24 | 0.253 | 0.24 | 0.25 | 0.26 | 0.20213 | 0.224 | 0.23739 | 0.2619 |  | 61
| 1 | 6 | 0.2457 | 0.21094 | 0.24 | 0.26 | 0.24 | 0.23 | 0.25 | 0.20485 | 0.2262 | 0.22984 | 0.266
| Mean |  | 0.25547 | 0.21028 | 0.24 | 0.25283 | 0.24 | 0.24 | 0.24833 | 0.20599 | 0.22432 | 0.23241 | 0.26345
| Median |  | 0.25665 | 0.21047 | 0.24 | 0.252 | 0.24 | 0.24 | 0.25 | 0.206 | 0.22455 | 0.23285 | 0.26345
| Std Dev. |  | 0.0053541 | 0.00087441 | 0 | 0.0040702 | 0 | 0.0063246 | 0.0075277 | 0.0033876 | 0.0023668 | 0.0033235 | 0.0035512
| Rel.Std.Dev. |  | 2.096% | 0.4158% | 0% | 1.61% | 0% | 2.635% | 3.031% | 1.645% | 1.055% | 1.43% | 1.348%
| PDM3 |  | 7.541% | -11.48% | 1.031% | 6.433% | 1.031% | 1.031% | 4.539% | -13.29% | -5.572% | -2.165% | 10.9%
| Z-Score (Absolute) |  | 0.81 | 1.56 | 0 | 0.67 | 0 | 0 | 0.44 | 1.79 | 0.82 | 0.4 | 1.23
| Analytical results for Sr in OREAS 143 (Certified Value 131 ppm) | [13 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 13 | 14 | [5 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 126 | 50.4 | 130 | 137 | 131.6128 | 118 | 135.0514 | [5 empty] | 1
| 1 | 2 | 130 | 50.6 | 130 | 131 | 131.6728 | 113 | 133.4143 | [5 empty] | #N/A
| 1 | 3 | 129 | 50.4 | 130 | 134 | 132.1772 | 122 | 129.0537 | [5 empty] | 16
| 1 | 4 | 133 | 49.5 | 130 | 130 | 130.7946 | 118 | 140.2861 | [5 empty] | 131.2679
| 1 | 5 | 129 | 49.6 | 130 | 133 | 130.6252 | 119 | 136.2548 | [5 empty] | 62
| 1 | 6 | 128 | 49.6 | 130 | 126 | 130.7628 | 119 | 130.3314
| Mean |  | 129.1667 | 50.0167 | 130 | 131.8333 | 131.2742 | 118.1667 | 134.0653
| Median |  | 129 | 50 | 130 | 132 | 131.2037 | 118.5 | 134.2328
| Std Dev. |  | 2.3166 | 0.49967 | 0 | 3.7639 | 0.63272 | 2.9269 | 4.0975
| Rel.Std.Dev. |  | 1.794% | 0.999% | 0% | 2.855% | 0.482% | 2.477% | 3.056%
| PDM3 |  | -1.601% | -61.9% | -0.9659% | 0.4307% | 0.004823% | -9.981% | 2.131%
| Z-Score (Absolute) |  | 0.31 | 29.42 | 0 | 0.67 | 0.47 | 4.35 | 1.5
| Analytical results for Ta in OREAS 143 (Certified Value 1 ppm) | [13 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 14 | 17 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | BF*OES/MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 0.9 | 1.1 | 0.8 | 0.9 | 1.1948 | 1.3 | [6 empty] | 1
| 1 | 2 | 0.7 | 1.1 | 0.7 | 0.7 | 1.2498 | 1.7 | [6 empty] | 3
| 1 | 3 | 1 | 1.1 | 0.7 | 1 | 1.184 | 1.2 | [6 empty] | 16
| 1 | 4 | 0.7 | 1.1 | 0.7 | 0.8 | 1.0612 | 1.1 | [6 empty] | 0.99606
| 1 | 5 | 0.9 | 1.1 | 0.7 | 0.7 | 1.1364 | 1.1 | [6 empty] | 63
| 1 | 6 | 0.8 | 1.1 | 0.7 | 0.9 | 1.2321 | 1.5
| Mean |  | 0.83333 | 1.1 | 0.71667 | 0.83333 | 1.1764 | 1.3167
| Median |  | 0.85 | 1.1 | 0.7 | 0.85 | 1.1894 | 1.25
| Std Dev. |  | 0.12111 | 2.4324e-16 | 0.040825 | 0.12111 | 0.06893 | 0.24014
| Rel.Std.Dev. |  | 14.53% | 2.211e-14% | 5.696% | 14.53% | 5.86% | 18.24%
| PDM3 |  | -16.34% | 10.43% | -28.05% | -16.34% | 18.1% | 32.19%
| Z-Score (Absolute) |  | 0.52 | 0.52 | 0.98 | 0.52 | 0.82 | 1.38
| Analytical results for Tb in OREAS 143 (Certified Value 0.8 ppm) | [13 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 0.9 | 0.7 | 0.83 | 0.9 | 0.6288 | 0.91749 | < 2 | [5 empty] | 1
| 1 | 2 | 0.9 | 0.7 | 0.83 | 0.8 | 0.6063 | 0.94627 | < 2 | [5 empty] | #N/A
| 1 | 3 | 0.9 | 0.7 | 0.88 | 0.8 | 0.6642 | 0.87548 | < 2 | [5 empty] | 16
| 1 | 4 | 0.9 | 0.7 | 0.85 | 0.8 | 0.6366 | 0.88654 | < 2 | [5 empty] | 0.80155
| 1 | 5 | 0.9 | 0.7 | 0.85 | 0.8 | 0.5997 | 0.90275 | < 2 | [5 empty] | 64
| 1 | 6 | 0.9 | 0.7 | 0.84 | 0.8 | 0.6554 | 0.95622 | < 2
| Mean |  | 0.9 | 0.7 | 0.84667 | 0.81667 | 0.63183 | 0.91413
| Median |  | 0.9 | 0.7 | 0.845 | 0.8 | 0.6327 | 0.91012
| Std Dev. |  | 0 | 1.2162e-16 | 0.018619 | 0.040825 | 0.025766 | 0.032248
| Rel.Std.Dev. |  | 0% | 1.737e-14% | 2.199% | 4.999% | 4.078% | 3.528%
| PDM3 |  | 12.28% | -12.67% | 5.629% | 1.886% | -21.17% | 14.05%
| Z-Score (Absolute) |  | 0.53 | 1.47 | 0 | 0.3 | 2.15 | 0.67 | 1.53
| Analytical results for Te in OREAS 143 (Indicative Value < 1 ppm) | [13 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 04 | 05 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 1 | < 6 | [10 empty] | 1
| 1 | 2 | < 1 | < 6 | [10 empty] | 4
| 1 | 3 | < 1 | < 6 | [10 empty] | 16
| 1 | 4 | < 1 | < 6 | [10 empty] | < 1
| 1 | 5 | < 1 | < 6 | [10 empty] | 21
| 1 | 6 | < 1 | < 6
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 143 (Certified Value 13 ppm) | [13 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 14 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | BF*OES/MS | [7 empty] | 1
2
| 1 | 1 | 13.4 | 11.5 | 13.1 | 13.2 | 13.9889 | [7 empty] | 1
| 1 | 2 | 13.6 | 11.6 | 13.2 | 12.9 | 13.6854 | [7 empty] | #N/A
| 1 | 3 | 13.5 | 11.4 | 13 | 13.5 | 13.5778 | [7 empty] | 16
| 1 | 4 | 13.2 | 11.3 | 13 | 13.3 | 13.7476 | [7 empty] | 13.0093
| 1 | 5 | 13.4 | 11.8 | 13.1 | 13.4 | 13.97 | [7 empty] | 65
| 1 | 6 | 13.2 | 11.8 | 13.2 | 12.9 | 13.8105
| Mean |  | 13.3833 | 11.5667 | 13.1 | 13.2 | 13.7967
| Median |  | 13.4 | 11.55 | 13.1 | 13.25 | 13.7791
| Std Dev. |  | 0.16021 | 0.20656 | 0.089443 | 0.25298 | 0.16114
| Rel.Std.Dev. |  | 1.197% | 1.786% | 0.6828% | 1.917% | 1.168%
| PDM3 |  | 2.875% | -11.09% | 0.6969% | 1.466% | 6.052%
| Z-Score (Absolute) |  | 0.67 | 6.01 | 0.37 | 0 | 2.19
| Analytical results for Ti in OREAS 143 (Certified Value 0.346 wt.%) | [13 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 12 | 13 | 14 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS |  |  | 3
3
| 1 | 1 | 0.36 | 0.19 | 0.34 | 0.35 | 0.348 | 0.348 | 0.342 | 0.33982 | 0.341 | 0.35442 |  |  | 1
| 1 | 2 | 0.35 | 0.19 | 0.34 | 0.35 | 0.336 | 0.36 | 0.348 | 0.34008 | 0.319 | 0.35651 |  |  | #N/A
| 1 | 3 | 0.36 | 0.19 | 0.35 | 0.36 | 0.33 | 0.33 | 0.348 | 0.3447 | 0.347 | 0.35876 |  |  | 16
| 1 | 4 | 0.37 | 0.19 | 0.34 | 0.34 | 0.336 | 0.348 | 0.342 | 0.3386 | 0.326 | 0.35939 |  |  | 0.34649
| 1 | 5 | 0.36 | 0.19 | 0.35 | 0.36 | 0.342 | 0.372 | 0.354 | 0.3468 | 0.332 | 0.35817 |  |  | 66
| 1 | 6 | 0.35 | 0.19 | 0.35 | 0.35 | 0.342 | 0.318 | 0.354 | 0.33394 | 0.331 | 0.36032
| Mean |  | 0.35833 | 0.19 | 0.345 | 0.35167 | 0.339 | 0.346 | 0.348 | 0.34066 | 0.33267 | 0.35793
| Median |  | 0.36 | 0.19 | 0.345 | 0.35 | 0.339 | 0.348 | 0.348 | 0.33995 | 0.3315 | 0.35846
| Std Dev. |  | 0.0075277 | 3.0405e-17 | 0.0054772 | 0.0075277 | 0.0062929 | 0.019596 | 0.0053666 | 0.0045681 | 0.010093 | 0.0021398
| Rel.Std.Dev. |  | 2.101% | 1.6e-14% | 1.588% | 2.141% | 1.856% | 5.664% | 1.542% | 1.341% | 3.034% | 0.5978%
| PDM3 |  | 3.42% | -45.16% | -0.4286% | 1.495% | -2.16% | -0.14% | 0.4372% | -1.682% | -3.988% | 3.303%
| Z-Score (Absolute) |  | 1.35 | 16.17 | 0.04 | 0.66 | 0.69 | 0.04 | 0.24 | 0.49 | 1.32 | 1.31
| Analytical results for Tl in OREAS 143 (Indicative Value 0.86 ppm) | [13 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 04 | 05 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 0.9 | < 1 | 0.9 | 0.7 | [8 empty] | 1
| 1 | 2 | 0.8 | < 1 | 0.9 | 0.7 | [8 empty] | 5
| 1 | 3 | 1 | < 1 | 0.9 | 0.8 | [8 empty] | 16
| 1 | 4 | 0.9 | < 1 | 0.9 | 0.8 | [8 empty] | 0.86111
| 1 | 5 | 0.9 | < 1 | 0.9 | 0.8 | [8 empty] | 22
| 1 | 6 | 0.9 | < 1 | 0.9 | 0.9
| Mean |  | 0.9 |  | 0.9 | 0.78333
| Median |  | 0.9 |  | 0.9 | 0.8
| Std Dev. |  | 0.063246 |  | 0 | 0.075277
| Rel.Std.Dev. |  | 7.027% |  | 0% | 9.61%
| PDM3 |  | 4.516% |  | 4.516% | -9.032%
| Z-Score (Absolute) |  | 0.67 | 3.95 | 0.67 | 0.67
| Analytical results for Tm in OREAS 143 (Certified Value 0.37 ppm) | [13 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 0.4 | < 1 | 0.39 | 0.4 | 0.274 | 0.4005 | < 2 | [5 empty] | 1
| 1 | 2 | 0.4 | < 1 | 0.41 | 0.5 | 0.2768 | 0.35324 | < 2 | [5 empty] | #N/A
| 1 | 3 | 0.4 | < 1 | 0.4 | 0.4 | 0.3037 | 0.35419 | < 2 | [5 empty] | 16
| 1 | 4 | 0.4 | < 1 | 0.38 | 0.5 | 0.2724 | 0.30689 | < 2 | [5 empty] | 0.37374
| 1 | 5 | 0.4 | < 1 | 0.37 | 0.4 | 0.2853 | 0.36671 | < 2 | [5 empty] | 67
| 1 | 6 | 0.4 | < 1 | 0.4 | 0.4 | 0.2948 | 0.37361 | < 2
| Mean |  | 0.4 |  | 0.39167 | 0.43333 | 0.2845 | 0.35919
| Median |  | 0.4 |  | 0.395 | 0.4 | 0.28105 | 0.36045
| Std Dev. |  | 6.0809e-17 |  | 0.01472 | 0.05164 | 0.01258 | 0.030875
| Rel.Std.Dev. |  | 1.52e-14% |  | 3.758% | 11.92% | 4.422% | 8.596%
| PDM3 |  | 7.027% |  | 4.797% | 15.95% | -23.88% | -3.892%
| Z-Score (Absolute) |  | 0 | 1.65 | 0.14 | 0.55 | 1.91 | 0.67 | 9.91
| Analytical results for U in OREAS 143 (Certified Value 3.08 ppm) | [13 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 3.3 | 2.5 | 3.2 | 3.1 | 2.2681 | 3.3528 | [6 empty] | 1
| 1 | 2 | 3.8 | 2.8 | 3.2 | 3.2 | 2.2752 | 3.054 | [6 empty] | #N/A
| 1 | 3 | 3.2 | 2.7 | 3.2 | 3.1 | 2.3785 | 2.9142 | [6 empty] | 16
| 1 | 4 | 3.4 | 2.7 | 3.2 | 3.1 | 2.3362 | 2.7352 | [6 empty] | 3.0795
| 1 | 5 | 3.5 | 2.8 | 3.1 | 3.1 | 2.0768 | 3.6288 | [6 empty] | 68
| 1 | 6 | 3.1 | 2.5 | 3.1 | 3.4 | 2.3778 | 2.68
| Mean |  | 3.3833 | 2.6667 | 3.1667 | 3.1667 | 2.2854 | 3.0608
| Median |  | 3.35 | 2.7 | 3.2 | 3.1 | 2.3057 | 2.9841
| Std Dev. |  | 0.24833 | 0.13663 | 0.05164 | 0.12111 | 0.11287 | 0.36891
| Rel.Std.Dev. |  | 7.34% | 5.123% | 1.631% | 3.824% | 4.939% | 12.05%
| PDM3 |  | 9.866% | -13.41% | 2.831% | 2.831% | -25.79% | -0.606%
| Z-Score (Absolute) |  | 1.13 | 1.87 | 0.22 | 0.22 | 3.46 | 0.22
| Analytical results for V in OREAS 143 (Certified Value 70 ppm) | [13 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 13 | [6 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | [6 empty] | 0
0
| 1 | 1 | 122 | 55.8 | 70 | 82 | 75.79 | 71 | [6 empty] | 1
| 1 | 2 | 76 | 54.9 | 70 | 81 | 75.75 | 68 | [6 empty] | #N/A
| 1 | 3 | 123 | 56.6 | 70 | 75 | 75.19 | 70 | [6 empty] | 16
| 1 | 4 | 133 | 56.5 | 70 | 76 | 75.33 | 70 | [6 empty] | 69.9663
| 1 | 5 | 127 | 54.8 | 70 | 77 | 76.11 | 69 | [6 empty] | 69
| 1 | 6 | 120 | 55.6 | 70 | 79 | 76.62 | 72
| Mean |  | 116.8333 | 55.7 | 70 | 78.3333 | 75.7983 | 70
| Median |  | 122.5 | 55.7 | 70 | 78 | 75.77 | 70
| Std Dev. |  | 20.5272 | 0.7642 | 0 | 2.8048 | 0.52232 | 1.4142
| Rel.Std.Dev. |  | 17.57% | 1.372% | 0% | 3.581% | 0.6891% | 2.02%
| PDM3 |  | 66.99% | -20.39% | 0.04812% | 11.96% | 8.335% | 0.04812%
| Z-Score (Absolute) |  | 7.11 | 2.78 | 0.47 | 0.88 | 0.47 | 0.47
| Analytical results for W in OREAS 143 (Indicative Value 26.7 ppm) | [13 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 13 | 14 | [6 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | BF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 29 | < 5 | 24 | 26.6 | < 50 | 25.8039 | [6 empty] | 1
| 1 | 2 | 25 | < 5 | 24 | 28.2 | < 50 | 24.7616 | [6 empty] | 6
| 1 | 3 | 33 | < 5 | 25 | 29.6 | < 50 | 27.8638 | [6 empty] | 16
| 1 | 4 | 27 | < 5 | 24 | 26 | < 50 | 26.4413 | [6 empty] | 26.719
| 1 | 5 | 27 | < 5 | 24 | 27.8 | < 50 | 26.539 | [6 empty] | 23
| 1 | 6 | 30 | < 5 | 26 | 27.5 | < 50 | 26.1455
| Mean |  | 28.5 |  | 24.5 | 27.6167 |  | 26.2592
| Median |  | 28 |  | 24 | 27.65 |  | 26.2934
| Std Dev. |  | 2.8107 |  | 0.83666 | 1.2624 |  | 1.0148
| Rel.Std.Dev. |  | 9.862% |  | 3.415% | 4.571% |  | 3.865%
| PDM3 |  | 6.666% |  | -8.305% | 3.36% |  | -1.721%
| Z-Score (Absolute) |  | 1.24 | 10.01 | 0.49 | 0.86 | 0.27 | 0.27
| Analytical results for Y in OREAS 143 (Certified Value 25.6 ppm) | [13 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 13 | 14 | 17 |  |  |  |  | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | BF*OES/MS | PF*MS |  |  |  |  | 1
2
| 1 | 1 | 27 | 13.1 | 26.5 | 28 | 22.214 | 23 | 29.1852 | 25 |  |  |  |  | 1
| 1 | 2 | 26.3 | 12.8 | 26.1 | 27.6 | 22.569 | 21 | 29.7945 | 25 |  |  |  |  | 7
| 1 | 3 | 26.9 | 13 | 25.7 | 27.7 | 24.735 | 23 | 28.0784 | 26 |  |  |  |  | 16
| 1 | 4 | 26.4 | 13 | 25.9 | 26.1 | 21.199 | 24 | 27.6019 | 25 |  |  |  |  | 25.6143
| 1 | 5 | 26.9 | 12.9 | 26.4 | 29.2 | 21.455 | 22 | 29.0977 | 24 |  |  |  |  | 70
| 1 | 6 | 27.2 | 12.9 | 25.9 | 27 | 24.164 | 22 | 27.906 | 25
| Mean |  | 26.7833 | 12.95 | 26.0833 | 27.6 | 22.7227 | 22.5 | 28.6106 | 25
| Median |  | 26.9 | 12.95 | 26 | 27.65 | 22.3915 | 22.5 | 28.5881 | 25
| Std Dev. |  | 0.35449 | 0.10488 | 0.31252 | 1.0334 | 1.4379 | 1.0488 | 0.86789 | 0.63246
| Rel.Std.Dev. |  | 1.324% | 0.8099% | 1.198% | 3.744% | 6.328% | 4.661% | 3.033% | 2.53%
| PDM3 |  | 4.564% | -49.44% | 1.831% | 7.752% | -11.29% | -12.16% | 11.7% | -2.398%
| Z-Score (Absolute) |  | 0.34 | 3.48 | 0.15 | 0.57 | 0.78 | 0.84 | 0.85 | 0.15
| Analytical results for Yb in OREAS 143 (Certified Value 2.42 ppm) | [13 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 12 | 14 | 17 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 2.7 | 1.7 | 2.6 | 2.8 | 1.9264 | 2.8525 | 2 | [5 empty] | 1
| 1 | 2 | 2.6 | 1.6 | 2.5 | 2.7 | 1.9409 | 3.0579 | 3 | [5 empty] | #N/A
| 1 | 3 | 2.9 | 1.6 | 2.5 | 2.7 | 2.1191 | 2.7734 | 3 | [5 empty] | 16
| 1 | 4 | 2.8 | 1.7 | 2.3 | 2.7 | 1.9349 | 2.8018 | 2 | [5 empty] | 2.42
| 1 | 5 | 3.1 | 1.6 | 2.4 | 2.8 | 1.9232 | 3.0778 | 2 | [5 empty] | 71
| 1 | 6 | 2.8 | 1.7 | 2.4 | 2.5 | 2.0982 | 2.9129 | 3
| Mean |  | 2.8167 | 1.65 | 2.45 | 2.7 | 1.9905 | 2.9127 | 2.5
| Median |  | 2.8 | 1.65 | 2.45 | 2.7 | 1.9379 | 2.8827 | 2.5
| Std Dev. |  | 0.17224 | 0.054772 | 0.10488 | 0.10954 | 0.092007 | 0.12937 | 0.54772
| Rel.Std.Dev. |  | 6.115% | 3.32% | 4.281% | 4.057% | 4.622% | 4.442% | 21.91%
| PDM3 |  | 16.39% | -31.82% | 1.241% | 11.57% | -17.75% | 20.36% | 3.307%
| Z-Score (Absolute) |  | 0.56 | 2.15 | 0.29 | 0.29 | 1.36 | 0.79 | NA
Results from laboratory 17 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 143 (Certified Value 108 ppm) | [13 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 11 | 12 | 13 | 14 |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS |  | 0
0
| 1 | 1 | 106 | 89.4 | 100 | 120 | 100 | 120 | 180 | 150 | 93.159 | 128 | 100.0073 |  | 1
| 1 | 2 | 104 | 90 | 100 | 100 | 90 | 170 | 90 | < 100 | 93.0375 | 118 | 97.8158 |  | #N/A
| 1 | 3 | 106 | 90.3 | 100 | 110 | 90 | 130 | 130 | 110 | 92.315 | 192 | 100.0804 |  | 16
| 1 | 4 | 106 | 89.5 | 100 | 100 | 90 | 120 | 130 | 130 | 92.312 | 130 | 98.6869 |  | 107.5281
| 1 | 5 | 108 | 90.9 | 100 | 100 | 100 | 120 | 170 | < 100 | 92.9495 | 137 | 99.9297 |  | 72
| 1 | 6 | 103 | 89.6 | 100 | 130 | 90 | 120 | 300 | 140 | 92.7505 | 138 | 97.7402
| Mean |  | 105.5 | 89.95 | 100 | 110 | 93.3333 | 130 | 166.6667 | 132.5 | 92.7539 | 140.5 | 99.0434
| Median |  | 106 | 89.8 | 100 | 105 | 90 | 120 | 150 | 135 | 92.85 | 133.5 | 99.3083
| Std Dev. |  | 1.7607 | 0.57533 | 0 | 12.6491 | 5.164 | 20 | 72.8469 | 17.0783 | 0.36626 | 26.2431 | 1.1066
| Rel.Std.Dev. |  | 1.669% | 0.6396% | 0% | 11.5% | 5.533% | 15.38% | 43.71% | 12.89% | 0.3949% | 18.68% | 1.117%
| PDM3 |  | -1.886% | -16.35% | -7.001% | 2.299% | -13.2% | 20.9% | 55% | 23.22% | -13.74% | 30.66% | -7.891%
| Z-Score (Absolute) |  | 0.03 | 0.87 | 0.29 | 0.29 | 0.67 | 1.44 | 3.56 | 0 | 0.71 | 2.05 | 0.34
| Analytical results for Zr in OREAS 143 (Indicative Value 192 ppm) | [13 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 14 | 17 | [8 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*MS | [8 empty] | 0
0
| 1 | 1 | 192 | 9.3 | 204.0379 | 185 | [8 empty] | 1
| 1 | 2 | 187 | 9.1 | 197.7745 | 192 | [8 empty] | 18
| 1 | 3 | 189 | 8.5 | 196.7109 | 197 | [8 empty] | 16
| 1 | 4 | 191 | 9.3 | 198.7499 | 185 | [8 empty] | 191.6798
| 1 | 5 | 186 | 8.8 | 198.6186 | 173 | [8 empty] | 24
| 1 | 6 | 186 | 9.6 | 202.3437 | 189
| Mean |  | 188.5 | 9.1 | 199.7059 | 186.8333
| Median |  | 188 | 9.2 | 198.6843 | 187
| Std Dev. |  | 2.5884 | 0.39497 | 2.8468 | 8.1588
| Rel.Std.Dev. |  | 1.373% | 4.34% | 1.425% | 4.367%
| PDM3 |  | -1.659% | -95.25% | 4.187% | -2.528%
| Z-Score (Absolute) |  | 0.09 | 18.71 | 1.26 | 0.09
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C558: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C576: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C612: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C630: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C648: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C666: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C684: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C702: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C720: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C738: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C756: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C774: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C792: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C810: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C828: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C846: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C864: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C882: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1026: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 143 (Certified Value 0.497 ppm) | [17 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.4 | 0.48 | 0.4 | 0.6 | 0.49 | 0.48 | 0.52 | 0.5 | 0.54 | < 2 | 0.35 | 0.5027 | 0.51 | 0.9 |  |  | 1
| 1 | 2 | 0.6 | 0.49 | 0.4 | 0.7 | 0.49 | 0.49 | 0.48 | 0.49 | 0.53 | < 2 | 0.33 | 0.521 | 0.54 | 0.5 |  |  | 12
| 1 | 3 | 0.5 | 0.49 | 0.5 | 0.6 | 0.51 | 0.46 | 0.46 | 0.44 | 0.49 | < 2 | 0.35 | 0.52536 | 0.49 | 0.6 |  |  | 16
| 1 | 4 | 0.5 | 0.49 | 0.4 | 0.6 | 0.53 | 0.47 | 0.49 | 0.57 | 0.51 | < 2 | 0.39 | 0.51688 | 0.6 | 3.6 | 1.03 |  | 0.49727
| 1 | 5 | 0.6 | 0.5 | 0.4 | 0.6 | 0.45 | 0.49 | 0.55 | 0.48 | 0.52 | < 2 | 0.32 | 0.50892 | 0.45 | 0.6 |  |  | 74
| 1 | 6 | 0.8 | 0.5 | 0.4 | 0.7 | 0.48 | 0.49 | 0.47 | 0.48 | 0.53 | < 2 | 0.32 | 0.53233 | 0.53 | < 0.3
| Mean |  | 0.56667 | 0.49167 | 0.41667 | 0.63333 | 0.49167 | 0.48 | 0.495 | 0.49333 | 0.52 |  | 0.34333 | 0.51787 | 0.52 | 1.24 | 1.03
| Median |  | 0.55 | 0.49 | 0.4 | 0.6 | 0.49 | 0.485 | 0.485 | 0.485 | 0.525 |  | 0.34 | 0.51894 | 0.52 | 0.6 | 1.03
| Std Dev. |  | 0.13663 | 0.0075277 | 0.040825 | 0.05164 | 0.027142 | 0.012649 | 0.033912 | 0.04274 | 0.017889 |  | 0.026583 | 0.010828 | 0.050596 | 1.3278
| Rel.Std.Dev. |  | 24.11% | 1.531% | 9.798% | 8.154% | 5.52% | 2.635% | 6.851% | 8.663% | 3.44% |  | 7.743% | 2.091% | 9.73% | 107.1%
| PDM3 |  | 13.95% | -1.128% | -16.21% | 27.36% | -1.128% | -3.474% | -0.4575% | -0.7926% | 4.57% |  | -30.96% | 4.141% | 4.57% | 149.4% | 107.1%
| Z-Score (Absolute) |  | NA | 0.1 | NA | NA | 0.1 | 0.44 | 0 | 0.05 | 0.74 | 14.89 | 4.47 | 0.67 | 0.74 | NA | 15.78
| Analytical results for Al in OREAS 143 (Certified Value 6.37 wt.%) | [17 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 6.4 | 6.2 | 6.27 | 6.08 | 6.43 | 6.31 | 6.3 | 6.46 | 6.35 | 6.549 | 6.5505 | 7.18 | 6.5025 | 6.67 | 6.1994 | 5.9633 | 1
| 1 | 2 | 6.33 | 6.33 | 6.15 | 6.03 | 6.41 | 6.28 | 6.49 | 6.57 | 6.34 | 6.616 | 6.5356 | 7.3 | 6.6135 | 6.42 | 6.1824 | 6.006 | #N/A
| 1 | 3 | 6.52 | 6.29 | 6.34 | 6.21 | 6.42 | 6.3 | 6.37 | 6.56 | 6.21 | 6.61 | 6.5192 | 7.69 | 6.6767 | 6.59 | 6.1835 | 5.8974 | 16
| 1 | 4 | 6.43 | 6.21 | 6.28 | 6.17 | 6.47 | 6.19 | 6.29 | 6.66 | 6.21 | 6.665 | 6.5293 | 7.37 | 6.6164 | 6.5 | 6.2672 | 5.9841 | 6.3685
| 1 | 5 | 6.41 | 6.28 | 6.18 | 6.23 | 6.36 | 6.38 | 6.31 | 6.45 | 6.3 | 6.555 | 6.5313 | 7.13 | 6.6254 | 6.51 | 6.2062 | 5.9471 | 75
| 1 | 6 | 6.41 | 6.34 | 6.35 | 6.07 | 6.52 | 6.27 | 6.33 | 6.57 | 6.64 | 6.613 | 6.5723 | 7.21 | 6.812 | 6.64 | 6.1481 | 5.8934
| Mean |  | 6.4167 | 6.275 | 6.2617 | 6.1317 | 6.435 | 6.2883 | 6.3483 | 6.545 | 6.3417 | 6.6013 | 6.5397 | 7.3133 | 6.6411 | 6.555 | 6.1978 | 5.9486
| Median |  | 6.41 | 6.285 | 6.275 | 6.125 | 6.425 | 6.29 | 6.32 | 6.565 | 6.32 | 6.6115 | 6.5335 | 7.255 | 6.6209 | 6.55 | 6.1914 | 5.9552
| Std Dev. |  | 0.06121 | 0.058907 | 0.081833 | 0.082563 | 0.054681 | 0.061779 | 0.074944 | 0.078677 | 0.15842 | 0.043279 | 0.018962 | 0.20363 | 0.1013 | 0.094816 | 0.0395 | 0.045685
| Rel.Std.Dev. |  | 0.9539% | 0.9388% | 1.307% | 1.347% | 0.8497% | 0.9824% | 1.181% | 1.202% | 2.498% | 0.6556% | 0.29% | 2.784% | 1.525% | 1.446% | 0.6373% | 0.768%
| PDM3 |  | 0.757% | -1.467% | -1.677% | -3.718% | 1.045% | -1.258% | -0.316% | 2.772% | -0.4207% | 3.657% | 2.689% | 14.84% | 4.281% | 2.929% | -2.68% | -6.593%
| Z-Score (Absolute) |  | 0.14 | 0.45 | 0.51 | 1.06 | 0.22 | 0.4 | 0.14 | 0.69 | 0.17 | 0.92 | 0.66 | 3.93 | 1.09 | 0.73 | 0.78 | 1.83
| Analytical results for As in OREAS 143 (Certified Value 0.116 wt.%) | [17 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.1174 | 0.10705 | 0.114 | 0.12 | 0.1155 | 0.1155 | 0.115 | 0.1205 | 0.12 | 0.11796 | 0.1152 | 0.11443 | 0.0943 | 0.11426 | 0.10532 |  | 1
| 1 | 2 | 0.1171 | 0.10973 | 0.109 | 0.117 | 0.113 | 0.115 | 0.118 | 0.1195 | 0.1245 | 0.11806 | 0.1182 | 0.11549 | 0.113 | 0.11476 | 0.10648 |  | #N/A
| 1 | 3 | 0.1194 | 0.11064 | 0.113 | 0.117 | 0.116 | 0.115 | 0.116 | 0.1205 | 0.1155 | 0.1163 | 0.1194 | 0.11736 | 0.112 | 0.11634 | 0.10217 |  | 16
| 1 | 4 | 0.1198 | 0.11318 | 0.112 | 0.118 | 0.117 | 0.115 | 0.117 | 0.1235 | 0.117 | 0.11735 | 0.1198 | 0.12101 | 0.11 | 0.11717 | 0.10574 |  | 0.116
| 1 | 5 | 0.12 | 0.11257 | 0.11 | 0.114 | 0.116 | 0.117 | 0.1155 | 0.1175 | 0.1185 | 0.11578 | 0.1161 | 0.11449 | 0.106 | 0.11535 | 0.10411 |  | 76
| 1 | 6 | 0.1208 | 0.1134 | 0.113 | 0.117 | 0.1165 | 0.115 | 0.1145 | 0.12 | 0.124 | 0.11808 | 0.1178 | 0.11693 | 0.111 | 0.11502 | 0.10258
| Mean |  | 0.11908 | 0.1111 | 0.11183 | 0.11717 | 0.11567 | 0.11542 | 0.116 | 0.12025 | 0.11992 | 0.11726 | 0.11775 | 0.11662 | 0.10772 | 0.11548 | 0.1044
| Median |  | 0.1196 | 0.11161 | 0.1125 | 0.117 | 0.116 | 0.115 | 0.11575 | 0.12025 | 0.11925 | 0.11766 | 0.118 | 0.11621 | 0.1105 | 0.11519 | 0.10471
| Std Dev. |  | 0.0014945 | 0.0024617 | 0.0019408 | 0.0019408 | 0.0014024 | 0.00080104 | 0.0013038 | 0.0019429 | 0.0036799 | 0.00099316 | 0.0018086 | 0.0024695 | 0.007003 | 0.0010799 | 0.0017514
| Rel.Std.Dev. |  | 1.255% | 2.216% | 1.735% | 1.656% | 1.212% | 0.694% | 1.124% | 1.616% | 3.069% | 0.847% | 1.536% | 2.118% | 6.501% | 0.9351% | 1.678%
| PDM3 |  | 2.662% | -4.224% | -3.588% | 1.01% | -0.2834% | -0.4989% | 0.003994% | 3.668% | 3.381% | 1.087% | 1.513% | 0.5356% | -7.137% | -0.4414% | -9.997%
| Z-Score (Absolute) |  | 1.19 | 1.89 | 1.61 | 0.45 | 0.13 | 0.22 | 0 | 1.64 | 1.51 | 0.48 | 0.67 | 0.24 | 3.19 | 0.2 | 4.47
| Analytical results for B in OREAS 143 (Indicative Value 7.21 ppm) | [17 empty] | Indicative
B | Round | Replicate | Lab | [15 empty] | 1
| No | No | 18 | [15 empty] | ppm
|  |  | 4A*OES | [15 empty] | 2
2
| 1 | 1 | 10.61 | [15 empty] | 1
| 1 | 2 | 8.33 | [15 empty] | 13
| 1 | 3 | 4.24 | [15 empty] | 16
| 1 | 4 | 6.18 | [15 empty] | 7.2133
| 1 | 5 | 5.79 | [15 empty] | 26
| 1 | 6 | 8.13
| Mean |  | 7.2133
| Median |  | 7.155
| Std Dev. |  | 2.2609
| Rel.Std.Dev. |  | 31.34%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 143 (Certified Value 656 ppm) | [17 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 670 | 702 | 663 | 270 | 660 | 660 | 640 | 650 | 680 | 632.477 | 666.2 | 618 | 643.6074 | 670 | 603.71 |  | 1
| 1 | 2 | 669 | 712 | 655 | 524 | 660 | 660 | 660 | 660 | 670 | 629.046 | 665.06 | 619 | 655.4181 | 670 | 626.13 |  | #N/A
| 1 | 3 | 690 | 695 | 659 | 467 | 660 | 660 | 650 | 660 | 660 | 634.234 | 665.38 | 656 | 657.2868 | 720 | 615.06 |  | 16
| 1 | 4 | 686 | 688 | 646 | 427 | 670 | 650 | 640 | 670 | 670 | 634.857 | 665.55 | 631 | 671.1644 | 660 | 634.2 |  | 655.7836
| 1 | 5 | 687 | 693 | 648 | 525 | 660 | 670 | 640 | 650 | 670 | 630.597 | 671.76 | 613 | 651.6217 | 640 | 615.24 |  | 77
| 1 | 6 | 684 | 688 | 660 | 616 | 670 | 660 | 650 | 650 | 710 | 627.692 | 663.52 | 622 | 658.9444 | 640 | 618.33
| Mean |  | 681 | 696.3333 | 655.1667 | 471.5 | 663.3333 | 660 | 646.6667 | 656.6667 | 676.6667 | 631.4838 | 666.245 | 626.5 | 656.3405 | 666.6667 | 618.7783
| Median |  | 685 | 694 | 657 | 495.5 | 660 | 660 | 645 | 655 | 670 | 631.537 | 665.465 | 620.5 | 656.3524 | 665 | 616.785
| Std Dev. |  | 9.1214 | 9.2664 | 6.8532 | 117.5598 | 5.164 | 6.3246 | 8.165 | 8.165 | 17.5119 | 2.8648 | 2.8453 | 15.6301 | 9.0849 | 29.4392 | 10.4449
| Rel.Std.Dev. |  | 1.339% | 1.331% | 1.046% | 24.93% | 0.7785% | 0.9583% | 1.263% | 1.243% | 2.588% | 0.4537% | 0.4271% | 2.495% | 1.384% | 4.416% | 1.688%
| PDM3 |  | 3.845% | 6.183% | -0.09407% | -28.1% | 1.151% | 0.643% | -1.39% | 0.1347% | 3.184% | -3.705% | 1.595% | -4.465% | 0.08492% | 1.66% | -5.643%
| Z-Score (Absolute) |  | 1.64 | 2.67 | 0.1 | 12.49 | 0.45 | 0.22 | 0.67 | 0 | 1.35 | 1.7 | 0.65 | 2.03 | 0.02 | 0.67 | 2.55
| Analytical results for Be in OREAS 143 (Certified Value 3.05 ppm) | [17 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 3.6 | 2.59 | 2.9 | 3 | 2.89 | 3.07 | 2.86 | 3.1 | 3 | 3.252 | 2.7839 | 3 | 2.9999 | 3.4 | 3.14 |  | 1
| 1 | 2 | 2.8 | 2.52 | 3 | 3.2 | 2.88 | 3.05 | 2.98 | 2.97 | 3.16 | 3.038 | 2.8444 | 3 | 3.206 | 3.3 | 3.07 |  | #N/A
| 1 | 3 | 3.1 | 2.6 | 2.9 | 3.3 | 3.07 | 3.08 | 2.91 | 3.03 | 2.9 | 3.332 | 2.8244 | 3 | 3.0483 | 3.5 | 2.92 |  | 16
| 1 | 4 | 3.2 | 2.43 | 2.9 | 2.9 | 2.92 | 2.9 | 3.05 | 3.23 | 2.95 | 3.312 | 2.8626 | 3 | 3.0241 | 3.3 | 3.13 |  | 3.0464
| 1 | 5 | 3.2 | 2.56 | 3 | 3.5 | 2.88 | 3.14 | 2.93 | 3.01 | 2.94 | 3.385 | 2.6394 | 3 | 3.0848 | 3.2 | 3.05 |  | 78
| 1 | 6 | 3.1 | 2.56 | 2.9 | 2.8 | 2.85 | 2.97 | 2.97 | 3.1 | 3.24 | 2.993 | 2.5684 | 3 | 3.0282 | 3.1 | 2.95
| Mean |  | 3.1667 | 2.5433 | 2.9333 | 3.1167 | 2.915 | 3.035 | 2.95 | 3.0733 | 3.0317 | 3.2187 | 2.7538 | 3 | 3.0652 | 3.3 | 3.0433
| Median |  | 3.15 | 2.56 | 2.9 | 3.1 | 2.885 | 3.06 | 2.95 | 3.065 | 2.975 | 3.282 | 2.8041 | 3 | 3.0382 | 3.3 | 3.06
| Std Dev. |  | 0.2582 | 0.062183 | 0.05164 | 0.26394 | 0.079183 | 0.085965 | 0.065422 | 0.092232 | 0.13659 | 0.16365 | 0.12116 | 0 | 0.074593 | 0.14142 | 0.091141
| Rel.Std.Dev. |  | 8.154% | 2.445% | 1.76% | 8.469% | 2.716% | 2.832% | 2.218% | 3.001% | 4.505% | 5.084% | 4.4% | 0% | 2.434% | 4.285% | 2.995%
| PDM3 |  | 3.949% | -16.51% | -3.71% | 2.308% | -4.312% | -0.373% | -3.163% | 0.8853% | -0.4824% | 5.656% | -9.603% | -1.522% | 0.6192% | 8.326% | -0.09946%
| Z-Score (Absolute) |  | 0.88 | 3.43 | 0.73 | 0.54 | 0.86 | 0.03 | 0.62 | 0.24 | 0.05 | 1.24 | 1.97 | NA | 0.18 | 1.8 | 0.03
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 143 (Certified Value 23.4 ppm) | [17 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 22.7 | 28.6 | 20.4 | 23.2 | 20.8 | 23.8 | 23.8 | 24.3 | 23.6 | 31.22 | 25.1 | 21.4 | 24.0048 | 23.7 | 23.59 | 15.72 | 1
| 1 | 2 | 23.3 | 29.31 | 21 | 23.2 | 22.7 | 23.6 | 25.1 | 26 | 24.3 | 25.519 | 24.63 | 22.9 | 24.6469 | 23.2 | 22.87 | 18.47 | 1
| 1 | 3 | 23.9 | 27.85 | 21.2 | 23.9 | 23 | 26.9 | 24.2 | 25.3 | 22.4 | 24.988 | 23.66 | 23.4 | 24.9085 | 22.7 | 21.63 | 16.56 | 16
| 1 | 4 | 24.3 | 28.2 | 20.5 | 24.8 | 22.1 | 22 | 27.7 | 25.3 | 24.3 | 26.495 | 25.69 | 21.8 | 24.9373 | 22 | 23.54 | 18.71 | 23.3702
| 1 | 5 | 24.6 | 27.46 | 21.2 | 22.5 | 21.9 | 22.3 | 24.4 | 24.1 | 22.2 | 30.465 | 25.39 | 21.6 | 24.3197 | 22 | 21.38 | 15 | 79
| 1 | 6 | 23.1 | 27.42 | 21.9 | 24.3 | 21.4 | 26.1 | 22.9 | 24.3 | 23.4 | 30.524 | 23.76 | 21.4 | 24.1748 | 25 | 22.96 | 21.91
| Mean |  | 23.65 | 28.14 | 21.0333 | 23.65 | 21.9833 | 24.1167 | 24.6833 | 24.8833 | 23.3667 | 28.2018 | 24.705 | 22.0833 | 24.4987 | 23.1 | 22.6617 | 17.7283
| Median |  | 23.6 | 28.025 | 21.1 | 23.55 | 22 | 23.7 | 24.3 | 24.8 | 23.5 | 28.48 | 24.865 | 21.7 | 24.4833 | 22.95 | 22.915 | 17.515
| Std Dev. |  | 0.73689 | 0.72787 | 0.5465 | 0.84083 | 0.81343 | 1.9914 | 1.6461 | 0.76004 | 0.9048 | 2.8307 | 0.84689 | 0.85421 | 0.39068 | 1.1454 | 0.9458 | 2.5236
| Rel.Std.Dev. |  | 3.116% | 2.587% | 2.598% | 3.555% | 3.7% | 8.257% | 6.669% | 3.054% | 3.872% | 10.04% | 3.428% | 3.868% | 1.595% | 4.959% | 4.174% | 14.23%
| PDM3 |  | 1.197% | 20.41% | -9.999% | 1.197% | -5.934% | 3.194% | 5.619% | 6.475% | -0.01492% | 20.67% | 5.712% | -5.506% | 4.829% | -1.156% | -3.032% | -24.14%
| Z-Score (Absolute) |  | 0 | 2.9 | 1.69 | 0 | 1.08 | 0.3 | 0.67 | 0.8 | 0.18 | 2.94 | 0.68 | 1.01 | 0.55 | 0.36 | 0.64 | 3.82
| Analytical results for Ca in OREAS 143 (Certified Value 6.85 wt.%) | [17 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 6.94 | 6.87 | 6.54 | 6.69 | 6.82 | 7.01 | 6.77 | 6.92 | 6.97 | 7.101 | 6.6787 | 7.38 | 6.8624 | 6.9 | 6.8219 | 6.3711 | 1
| 1 | 2 | 6.86 | 7.01 | 6.6 | 6.32 | 6.83 | 6.78 | 6.91 | 7 | 6.78 | 7.15 | 6.6928 | 7.47 | 6.9009 | 6.68 | 6.7384 | 6.4745 | #N/A
| 1 | 3 | 6.91 | 6.97 | 6.79 | 6.6 | 6.82 | 6.82 | 6.93 | 6.96 | 6.76 | 7.03 | 6.687 | 7.71 | 7.0953 | 6.71 | 6.6286 | 6.2221 | 16
| 1 | 4 | 6.93 | 6.94 | 6.73 | 6.73 | 6.86 | 6.79 | 6.73 | 7.54 | 6.77 | 6.993 | 6.6858 | 7.36 | 7.1616 | 6.59 | 6.7885 | 6.5921 | 6.8532
| 1 | 5 | 6.92 | 7.01 | 6.62 | 6.78 | 7 | 6.97 | 6.62 | 7.02 | 6.91 | 6.933 | 6.6825 | 7.46 | 7.0829 | 6.57 | 6.636 | 6.3511 | 80
| 1 | 6 | 7.1 | 6.98 | 6.76 | 6.56 | 7.05 | 6.96 | 6.68 | 7.14 | 7.23 | 7.047 | 6.686 | 7.49 | 7.1677 | 6.92 | 6.7211 | 6.4944
| Mean |  | 6.9433 | 6.9633 | 6.6733 | 6.6133 | 6.8967 | 6.8883 | 6.7733 | 7.0967 | 6.9033 | 7.0423 | 6.6855 | 7.4783 | 7.0452 | 6.7283 | 6.7224 | 6.4176
| Median |  | 6.925 | 6.975 | 6.675 | 6.645 | 6.845 | 6.89 | 6.75 | 7.01 | 6.845 | 7.0385 | 6.6859 | 7.465 | 7.0891 | 6.695 | 6.7298 | 6.4228
| Std Dev. |  | 0.08165 | 0.052789 | 0.10033 | 0.16513 | 0.10172 | 0.10265 | 0.12437 | 0.22958 | 0.18151 | 0.076935 | 0.0047141 | 0.12481 | 0.13168 | 0.15039 | 0.078474 | 0.12989
| Rel.Std.Dev. |  | 1.176% | 0.7581% | 1.503% | 2.497% | 1.475% | 1.49% | 1.836% | 3.235% | 2.629% | 1.092% | 0.07051% | 1.669% | 1.869% | 2.235% | 1.167% | 2.024%
| PDM3 |  | 1.315% | 1.606% | -2.625% | -3.501% | 0.6337% | 0.5121% | -1.166% | 3.552% | 0.7309% | 2.759% | -2.448% | 9.121% | 2.8% | -1.823% | -1.909% | -6.357%
| Z-Score (Absolute) |  | 0.22 | 0.3 | 0.93 | 1.19 | 0.02 | 0.02 | 0.51 | 0.87 | 0.05 | 0.64 | 0.88 | 2.49 | 0.65 | 0.7 | 0.72 | 2.02
| Analytical results for Cd in OREAS 143 (Indicative Value 0.46 ppm) | [17 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 2
2
| 1 | 1 | 0.3 | 0.2 | 0.23 | 0.5 | 1.07 | 0.25 | < 0.02 | 0.35 | < 0.02 | < 1 | 2.1828 | < 0.02 | 0.27827 | 3.9 | 0.09 | 1.17 | 1
| 1 | 2 | 0.3 | 0.2 | 0.21 | 0.7 | 0.18 | 0.02 | 0.34 | 0.32 | 2.36 | < 1 | 2.2236 | < 0.02 | 0.31084 | 4 | 0.1 | 1.1 | 2
| 1 | 3 | 0.2 | 0.2 | 0.22 | 0.5 | 0.34 | 3.64 | 1.09 | 0.31 | < 0.02 | < 1 | 2.2848 | < 0.02 | 0.30376 | 4.6 | 0.12 | 0.788 | 16
| 1 | 4 |  | 0.21 | 0.23 | < 0.3 | 0.31 | 0.61 | 0.92 | 0.35 | < 0.02 | < 1 | 2.3256 | < 0.02 | 0.29537 | 3.8 | 0.22 | 0.77 | 0.46287
| 1 | 5 | 0.3 | 0.19 | 0.22 | 0.5 | 0.81 | 0.33 | < 0.02 | 0.4 | 0.92 | < 1 | 2.3766 | < 0.02 | 0.30586 | 3.9 | 0.14 | 0.87 | 27
| 1 | 6 | 0.3 | 0.19 | 0.23 | 0.8 | 0.21 | 0.24 | < 0.02 | 0.35 | 0.68 | < 1 | 2.2338 | < 0.02 | 0.28763 | 3.7 | 0.03 | 1
| Mean |  | 0.28 | 0.19833 | 0.22333 | 0.6 | 0.48667 | 0.84833 | 0.78333 | 0.34667 | 1.32 |  | 2.2712 |  | 0.29695 | 3.9833 | 0.11667 | 0.94967
| Median |  | 0.3 | 0.2 | 0.225 | 0.5 | 0.325 | 0.29 | 0.92 | 0.35 | 0.92 |  | 2.2593 |  | 0.29956 | 3.9 | 0.11 | 0.935
| Std Dev. |  | 0.044721 | 0.0075277 | 0.008165 | 0.14142 | 0.36555 | 1.3808 | 0.39323 | 0.031411 | 0.90863 |  | 0.071739 |  | 0.012302 | 0.31885 | 0.062823 | 0.1664
| Rel.Std.Dev. |  | 15.97% | 3.795% | 3.656% | 23.57% | 75.11% | 162.8% | 50.2% | 9.061% | 68.84% |  | 3.159% |  | 4.143% | 8.005% | 53.85% | 17.52%
| PDM3 |  | -39.51% | -57.15% | -51.75% | 29.62% | 5.14% | 83.27% | 69.23% | -25.11% | 185.2% |  | 390.7% |  | -35.85% | 760.6% | -74.8% | 105.2%
| Z-Score (Absolute) |  | NA | 0.74 | 0.67 | 0.1 | 0 | 0.93 | 0.23 | 0.36 | 0.46 | 0.03 | 4.57 | 1.22 | 0.49 | 8.95 | 0.95 | 1.19
| Analytical results for Ce in OREAS 143 (Certified Value 67 ppm) | [17 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
1
| 1 | 1 | 58.4 | 64.05 | 62.45 | 62.1 | 65.7 | 66.4 | 72.2 | 68.5 | 72.4 | 88.4238 | 90.9 | 65.4835 | 48.8 | 71.2311 |  |  | 1
| 1 | 2 | 64.8 | 62.99 | 64.2 | 64.4 | 67.4 | 66.7 | 74.6 | 71.4 | 71.5 | 88.7522 | 91.2 | 68.3861 | 47.4 | 73.7951 |  |  | 3
| 1 | 3 | 67.4 | 63.25 | 63.28 | 63.7 | 70.3 | 67.6 | 72.2 | 68.3 | 69.5 | 88.4049 | 95.4 | 69.0352 | 46.8 | 67.7226 |  |  | 16
| 1 | 4 | 61.7 | 63.52 | 61.49 | 65.1 | 70.7 | 66 | 72.8 | 70.4 | 69.9 | 88.6283 | 89.7 | 70.4338 | 45.7 | 68.7016 |  |  | 67.0281
| 1 | 5 | 58.8 | 63.48 | 63.73 | 60.6 | 64 | 69.6 | 69.3 | 69.6 | 71.3 | 88.6348 | 90.2 | 66.4777 | 45.9 | 67.8314 |  |  | 81
| 1 | 6 | 62.9 | 63.37 | 63.64 | 63.2 | 67 | 68 | 70 | 66 | 74 | 88.1656 | 91.2 | 68.4903 | 46.7 | 70.4184
| Mean |  | 62.3333 | 63.4433 | 63.1317 | 63.1833 | 67.5167 | 67.3833 | 71.85 | 69.0333 | 71.4333 | 88.5016 | 91.4333 | 68.0511 | 46.8833 | 69.95
| Median |  | 62.3 | 63.425 | 63.46 | 63.45 | 67.2 | 67.15 | 72.2 | 69.05 | 71.4 | 88.5261 | 91.05 | 68.4382 | 46.75 | 69.56
| Std Dev. |  | 3.4766 | 0.35314 | 0.99393 | 1.629 | 2.6011 | 1.3182 | 1.9305 | 1.8875 | 1.6513 | 0.21207 | 2.0324 | 1.7899 | 1.1268 | 2.3518
| Rel.Std.Dev. |  | 5.577% | 0.5566% | 1.574% | 2.578% | 3.853% | 1.956% | 2.687% | 2.734% | 2.312% | 0.2396% | 2.223% | 2.63% | 2.403% | 3.362%
| PDM3 |  | -7.004% | -5.348% | -5.813% | -5.736% | 0.7288% | 0.5299% | 7.194% | 2.992% | 6.572% | 32.04% | 36.41% | 1.526% | -30.05% | 4.359%
| Z-Score (Absolute) |  | 0.87 | 0.7 | 0.75 | 0.74 | 0.04 | 0.06 | 0.65 | 0.2 | 0.59 | 3.32 | 3.79 | 0.04 | 3.35 | 0.35
| Analytical results for Co in OREAS 143 (Certified Value 13 ppm) | [17 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 14 | 12.73 | 13 | 12 | 11.5 | 13 | 14.5 | 13.8 | 12.8 | 13.008 | 11.6831 | 14.1 | 12.4166 | 12.5 | 13.2 | 13.66 | 1
| 1 | 2 | 14 | 12.73 | 13 | 12 | 11.8 | 12.8 | 14.4 | 14.1 | 13.2 | 12.759 | 11.5615 | 14 | 12.5994 | 12.2 | 13.9 | 13.58 | #N/A
| 1 | 3 | 14 | 12.77 | 13 | 12 | 12.6 | 16.6 | 14.2 | 13.8 | 12.4 | 12.772 | 11.7091 | 14.9 | 12.9598 | 12.2 | 13.4 | 15.92 | 16
| 1 | 4 | 14 | 12.22 | 13 | 12 | 12.6 | 13.2 | 12.1 | 13.8 | 12.8 | 13.247 | 11.796 | 14.2 | 13.2529 | 11.6 | 13.7 | 14.04 | 13.0289
| 1 | 5 | 14 | 12.74 | 13 | 12 | 11.6 | 14 | 13.3 | 16 | 12.8 | 12.976 | 11.593 | 13.4 | 12.6114 | 11.9 | 13.4 | 15.19 | 82
| 1 | 6 | 14 | 13.02 | 13 | 12 | 12.1 | 13.2 | 13.8 | 13.4 | 13.4 | 13.167 | 11.3248 | 13.8 | 12.7684 | 12.4 | 13.2 | 13.38
| Mean |  | 14 | 12.7017 | 13 | 12 | 12.0333 | 13.8 | 13.7167 | 14.15 | 12.9 | 12.9882 | 11.6113 | 14.0667 | 12.7681 | 12.1333 | 13.4667 | 14.295
| Median |  | 14 | 12.735 | 13 | 12 | 11.95 | 13.2 | 14 | 13.8 | 12.8 | 12.992 | 11.6381 | 14.05 | 12.6899 | 12.2 | 13.4 | 13.85
| Std Dev. |  | 0 | 0.26118 | 0 | 0 | 0.48442 | 1.4311 | 0.90646 | 0.93327 | 0.35214 | 0.19938 | 0.16353 | 0.49666 | 0.29945 | 0.33267 | 0.28048 | 1.0255
| Rel.Std.Dev. |  | 0% | 2.056% | 0% | 0% | 4.026% | 10.37% | 6.608% | 6.596% | 2.73% | 1.535% | 1.408% | 3.531% | 2.345% | 2.742% | 2.083% | 7.174%
| PDM3 |  | 7.453% | -2.512% | -0.2219% | -7.897% | -7.641% | 5.918% | 5.279% | 8.605% | -0.9894% | -0.3127% | -10.88% | 7.965% | -2.002% | -6.874% | 3.36% | 9.718%
| Z-Score (Absolute) |  | NA | 0.24 | NA | NA | 0.79 | 0.67 | 0.61 | 0.97 | 0.07 | 0 | 1.14 | 0.9 | 0.18 | 0.71 | 0.4 | 1.09
Results from laboratories 1, 4 and 5 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 143 (Certified Value 68 ppm) | [17 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 91 | 81 | 59 | 61 | 64 | 68 | 67 | 68 | 64 | 74.146 | 53 | 65.3046 | 65 | 72.23 |  |  | 1
| 1 | 2 | 81 | 80 | 57 | 91 | 65 | 65 | 68 | 71 | 65 | 75.342 | 51 | 65.3257 | 62 | 72.2 |  |  | #N/A
| 1 | 3 | 85 | 77 | 56 | 83 | 65 | 66 | 66 | 71 | 66 | 72.744 | 57 | 67.0611 | 65 | 73.31 |  |  | 16
| 1 | 4 | 83 | 76 | 60 | 65 | 63 | 66 | 65 | 70 | 65 | 75.621 | 53 | 66.7654 | 62 | 76.69 |  |  | 67.9856
| 1 | 5 | 83 | 76 | 61 | 113 | 65 | 66 | 68 | 69 | 65 | 70.522 | 53 | 63.9418 | 60 | 66.399 |  |  | 83
| 1 | 6 | 81 | 78 | 56 | 71 | 64 | 66 | 65 | 70 | 66 | 74.813 | 53 | 64.1458 | 60 | 70.43
| Mean |  | 84 | 78 | 58.1667 | 80.6667 | 64.3333 | 66.1667 | 66.5 | 69.8333 | 65.1667 | 73.8647 | 53.3333 | 65.4241 | 62.3333 | 71.8765
| Median |  | 83 | 77.5 | 58 | 77 | 64.5 | 66 | 66.5 | 70 | 65 | 74.4795 | 53 | 65.3151 | 62 | 72.215
| Std Dev. |  | 3.7417 | 2.0976 | 2.137 | 19.4079 | 0.8165 | 0.98319 | 1.3784 | 1.169 | 0.75277 | 1.9326 | 1.9664 | 1.2911 | 2.2509 | 3.3926
| Rel.Std.Dev. |  | 4.454% | 2.689% | 3.674% | 24.06% | 1.269% | 1.486% | 2.073% | 1.674% | 1.155% | 2.616% | 3.687% | 1.973% | 3.611% | 4.72%
| PDM3 |  | 23.56% | 14.73% | -14.44% | 18.65% | -5.372% | -2.675% | -2.185% | 2.718% | -4.146% | 8.647% | -21.55% | -3.768% | -8.314% | 5.723%
| Z-Score (Absolute) |  | 2.5 | 1.65 | 1.15 | 2.03 | 0.28 | 0.02 | 0.02 | 0.49 | 0.16 | 1.06 | 1.84 | 0.13 | 0.57 | 0.78
| Analytical results for Cs in OREAS 143 (Certified Value 9.31 ppm) | [17 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 9.4 | 9.36 | 9.1 | 9.44 | 8.85 | 9.65 | 10.2 | 9.93 | 9.63 | 8.5421 | 11 | 9.1967 | 8.85 | 9.1702 |  |  | 1
| 1 | 2 | 9.4 | 9.08 | 8.8 | 9.62 | 9.06 | 9.66 | 10.4 | 10.45 | 9.76 | 8.4617 | 11 | 9.4312 | 8.81 | 8.9857 |  |  | 6
| 1 | 3 | 9.7 | 9.05 | 8.9 | 9.33 | 9.26 | 9.25 | 9.93 | 10.1 | 9.45 | 8.3618 | 12 | 9.5793 | 8.93 | 8.9895 |  |  | 16
| 1 | 4 | 8.9 | 8.96 | 9 | 9.62 | 9.54 | 9.8 | 9.86 | 10.25 | 9.51 | 8.5495 | 11 | 9.565 | 8.5 | 8.9286 |  |  | 9.3066
| 1 | 5 | 8.9 | 9.16 | 8.9 | 9.09 | 8.4 | 9.82 | 9.89 | 10.15 | 9.62 | 8.3426 | 11 | 9.2996 | 8.61 | 8.9943 |  |  | 84
| 1 | 6 | 9.5 | 8.9 | 8.9 | 9.58 | 8.87 | 9.66 | 10.25 | 9.96 | 10.1 | 8.428 | 11 | 9.2873 | 8.88 | 9.3693
| Mean |  | 9.3 | 9.085 | 8.9333 | 9.4467 | 8.9967 | 9.64 | 10.0883 | 10.14 | 9.6783 | 8.4476 | 11.1667 | 9.3932 | 8.7633 | 9.0729
| Median |  | 9.4 | 9.065 | 8.9 | 9.51 | 8.965 | 9.66 | 10.065 | 10.125 | 9.625 | 8.4449 | 11 | 9.3654 | 8.83 | 8.9919
| Std Dev. |  | 0.32863 | 0.1627 | 0.10328 | 0.20916 | 0.39042 | 0.20543 | 0.22463 | 0.19308 | 0.23267 | 0.087489 | 0.40825 | 0.1576 | 0.16967 | 0.16666
| Rel.Std.Dev. |  | 3.534% | 1.791% | 1.156% | 2.214% | 4.34% | 2.131% | 2.227% | 1.904% | 2.404% | 1.036% | 3.656% | 1.678% | 1.936% | 1.837%
| PDM3 |  | -0.07062% | -2.381% | -4.01% | 1.505% | -3.33% | 3.583% | 8.4% | 8.955% | 3.995% | -9.229% | 19.99% | 0.9309% | -5.837% | -2.51%
| Z-Score (Absolute) |  | 0.09 | 0.52 | 0.82 | 0.2 | 0.69 | 0.58 | 1.47 | 1.57 | 0.66 | 1.78 | 3.6 | 0.09 | 1.15 | 0.54
| Analytical results for Cu in OREAS 143 (Certified Value 328 ppm) | [17 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 323 | 325.93 | 324 | 323 | 336 | 335 | 321 | 338 | 325 | 337.366 | 324.65 | 329 | 318.9936 | 334 | 334.3 | 328.3584 | 1
| 1 | 2 | 316 | 324.87 | 330 | 327 | 334 | 331 | 329 | 340 | 337 | 336.516 | 328.05 | 329 | 325.2435 | 327 | 333.6 | 320.997 | #N/A
| 1 | 3 | 327 | 319.27 | 329 | 323 | 336 | 332 | 323 | 331 | 318 | 338.677 | 332.02 | 326 | 323.5446 | 331 | 335.9 | 302.0402 | 16
| 1 | 4 | 327 | 315.41 | 326 | 323 | 338 | 328 | 330 | 346 | 315 | 338.749 | 327.76 | 325 | 331.7325 | 324 | 337.9 | 328.1305 | 327.7761
| 1 | 5 | 340 | 320.17 | 324 | 322 | 337 | 338 | 321 | 352 | 324 | 334.443 | 325.33 | 326 | 324.496 | 325 | 332.5 | 307.4171 | 85
| 1 | 6 | 325 | 317.86 | 327 | 319 | 337 | 331 | 321 | 339 | 336 | 340.01 | 321.77 | 331 | 329.7568 | 341 | 331.3 | 311.54
| Mean |  | 326.3333 | 320.585 | 326.6667 | 322.8333 | 336.3333 | 332.5 | 324.1667 | 341 | 325.8333 | 337.6268 | 326.5967 | 327.6667 | 325.6278 | 330.3333 | 334.25 | 316.4139
| Median |  | 326 | 319.72 | 326.5 | 323 | 336.5 | 331.5 | 322 | 339.5 | 324.5 | 338.0215 | 326.545 | 327.5 | 324.8698 | 329 | 333.95 | 316.2685
| Std Dev. |  | 7.8401 | 4.0752 | 2.5033 | 2.5626 | 1.3663 | 3.5071 | 4.2151 | 7.2111 | 9.0646 | 1.9738 | 3.5091 | 2.3381 | 4.5628 | 6.4395 | 2.3763 | 11.0658
| Rel.Std.Dev. |  | 2.402% | 1.271% | 0.7663% | 0.7938% | 0.4062% | 1.055% | 1.3% | 2.115% | 2.782% | 0.5846% | 1.074% | 0.7136% | 1.401% | 1.949% | 0.7109% | 3.497%
| PDM3 |  | -0.4402% | -2.194% | -0.3385% | -1.508% | 2.611% | 1.441% | -1.101% | 4.034% | -0.5927% | 3.005% | -0.3598% | -0.03339% | -0.6554% | 0.7802% | 1.975% | -3.466%
| Z-Score (Absolute) |  | 0.05 | 1.09 | 0.01 | 0.68 | 1.74 | 1.06 | 0.44 | 2.58 | 0.14 | 1.98 | 0.01 | 0.19 | 0.18 | 0.67 | 1.37 | 1.84
| Analytical results for Dy in OREAS 143 (Certified Value 3.08 ppm) | [17 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 2.9 | 3.45 | 2.98 | 3.4 | 2.8753 | 3.2729 | 2.683 | [9 empty] | 1
| 1 | 2 | 2.9 | 3.26 | 2.94 | 3.4 | 2.9159 | 3.3788 | 2.6078 | [9 empty] | #N/A
| 1 | 3 | 3 | 3.35 | 3.02 | 3.5 | 2.9334 | 3.4298 | 2.6679 | [9 empty] | 16
| 1 | 4 | 3.2 | 3.25 | 2.81 | 3.5 | 2.9621 | 3.4024 | 2.6445 | [9 empty] | 3.0797
| 1 | 5 | 2.8 | 3.18 | 2.89 | 3.5 | 2.6487 | 3.1457 | 2.7758 | [9 empty] | 86
| 1 | 6 | 3.6 | 3.18 | 2.94 | 3.4 | 2.9169 | 3.264 | 2.8421
| Mean |  | 3.0667 | 3.2783 | 2.93 | 3.45 | 2.8754 | 3.3156 | 2.7035
| Median |  | 2.95 | 3.255 | 2.94 | 3.45 | 2.9164 | 3.3258 | 2.6754
| Std Dev. |  | 0.29439 | 0.10496 | 0.073212 | 0.054772 | 0.11457 | 0.1076 | 0.08805
| Rel.Std.Dev. |  | 9.6% | 3.202% | 2.499% | 1.588% | 3.984% | 3.245% | 3.257%
| PDM3 |  | -0.4244% | 6.448% | -4.862% | 12.02% | -6.636% | 7.658% | -12.22%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.44 | 1.22 | 0.61 | 0.79 | 1.16
| Analytical results for Er in OREAS 143 (Certified Value 1.46 ppm) | [17 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 1.4 | 1.64 | 1.36 | 1.6 | 1.2844 | 1.4597 | 1.2265 | [9 empty] | 1
| 1 | 2 | 1.4 | 1.68 | 1.3 | 1.6 | 1.2477 | 1.5378 | 1.2301 | [9 empty] | #N/A
| 1 | 3 | 1.6 | 1.68 | 1.3 | 1.9 | 1.3052 | 1.5294 | 1.2651 | [9 empty] | 16
| 1 | 4 | 1.5 | 1.68 | 1.2 | 1.8 | 1.2516 | 1.5612 | 1.2399 | [9 empty] | 1.4582
| 1 | 5 | 1.5 | 1.66 | 1.3 | 1.8 | 1.1942 | 1.4575 | 1.2828 | [9 empty] | 87
| 1 | 6 | 1.5 | 1.65 | 1.23 | 1.7 | 1.2155 | 1.6061 | 1.3682
| Mean |  | 1.4833 | 1.665 | 1.2817 | 1.7333 | 1.2498 | 1.5253 | 1.2688
| Median |  | 1.5 | 1.67 | 1.3 | 1.75 | 1.2496 | 1.5336 | 1.2525
| Std Dev. |  | 0.075277 | 0.017607 | 0.057417 | 0.12111 | 0.041336 | 0.058105 | 0.053315
| Rel.Std.Dev. |  | 5.075% | 1.057% | 4.48% | 6.987% | 3.307% | 3.809% | 4.202%
| PDM3 |  | 1.726% | 14.18% | -12.1% | 18.87% | -14.29% | 4.603% | -12.99%
| Z-Score (Absolute) |  | 0 | 0.61 | 0.67 | 0.84 | 0.78 | 0.14 | 0.72
| Analytical results for Eu in OREAS 143 (Certified Value 1.11 ppm) | [17 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 1.2 | 1.07 | 1.06 | 1.09 | 1.1433 | 1.1985 | 1.13 | [9 empty] | 1
| 1 | 2 | 1.1 | 1.06 | 1.11 | 1.18 | 1.1716 | 1.2035 | 1.0991 | [9 empty] | #N/A
| 1 | 3 | 1.1 | 1.09 | 1.05 | 1.17 | 1.1371 | 1.2559 | 1.0717 | [9 empty] | 16
| 1 | 4 | 1.1 | 1.06 | 1.04 | 1.14 | 1.1714 | 1.2003 | 1.1101 | [9 empty] | 1.1124
| 1 | 5 | 1.2 | 1.04 | 1.12 | 1.12 | 1.0619 | 1.2205 | 1.152 | [9 empty] | 88
| 1 | 6 | 1.1 | 1.11 | 1.08 | 1.11 | 1.105 | 1.264 | 1.1916
| Mean |  | 1.1333 | 1.0717 | 1.0767 | 1.135 | 1.1317 | 1.2238 | 1.1258
| Median |  | 1.1 | 1.065 | 1.07 | 1.13 | 1.1402 | 1.212 | 1.12
| Std Dev. |  | 0.05164 | 0.024833 | 0.03266 | 0.035071 | 0.042223 | 0.029177 | 0.042237
| Rel.Std.Dev. |  | 4.556% | 2.317% | 3.033% | 3.09% | 3.731% | 2.384% | 3.752%
| PDM3 |  | 1.886% | -3.658% | -3.208% | 2.036% | 1.741% | 10.02% | 1.204%
| Z-Score (Absolute) |  | 0.18 | 6.79 | 6.22 | 0.37 | 0 | 10.4 | 0.67
| Analytical results for Fe in OREAS 143 (Certified Value 4.79 wt.%) | [17 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 4.76 | 4.62 | 4.73 | 4.87 | 4.72 | 4.76 | 4.59 | 4.78 | 4.83 | 4.927 | 4.7803 | 5.2 | 4.7492 | 4.65 | 4.7516 | 4.4322 | 1
| 1 | 2 | 4.76 | 4.76 | 4.72 | 4.64 | 4.74 | 4.73 | 4.74 | 4.83 | 4.91 | 4.926 | 4.7171 | 5.23 | 4.7876 | 4.56 | 4.7509 | 4.4426 | #N/A
| 1 | 3 | 4.86 | 4.69 | 4.79 | 5.03 | 4.77 | 4.76 | 4.66 | 4.82 | 4.69 | 4.919 | 4.7931 | 5.61 | 4.8218 | 4.74 | 4.7787 | 4.2561 | 16
| 1 | 4 | 4.81 | 4.65 | 4.72 | 4.89 | 4.78 | 4.73 | 4.72 | 4.92 | 4.71 | 4.938 | 4.7792 | 5.32 | 4.8962 | 4.52 | 4.8508 | 4.51 | 4.7907
| 1 | 5 | 4.82 | 4.72 | 4.73 | 5.02 | 4.76 | 4.86 | 4.64 | 4.74 | 4.77 | 4.84 | 4.7954 | 5.17 | 4.8338 | 4.39 | 4.7836 | 4.3397 | 89
| 1 | 6 | 4.8 | 4.72 | 4.83 | 5.02 | 4.82 | 4.77 | 4.66 | 4.8 | 4.99 | 4.927 | 4.702 | 5.22 | 4.8697 | 4.32 | 4.7946 | 4.2475
| Mean |  | 4.8017 | 4.6933 | 4.7533 | 4.9117 | 4.765 | 4.7683 | 4.6683 | 4.815 | 4.8167 | 4.9128 | 4.7612 | 5.2917 | 4.8264 | 4.53 | 4.785 | 4.3714
| Median |  | 4.805 | 4.705 | 4.73 | 4.955 | 4.765 | 4.76 | 4.66 | 4.81 | 4.8 | 4.9265 | 4.7798 | 5.225 | 4.8278 | 4.54 | 4.7812 | 4.3859
| Std Dev. |  | 0.038166 | 0.051251 | 0.045898 | 0.15065 | 0.034496 | 0.047924 | 0.054559 | 0.060581 | 0.11708 | 0.036196 | 0.040814 | 0.16388 | 0.053432 | 0.15697 | 0.03672 | 0.10735
| Rel.Std.Dev. |  | 0.7949% | 1.092% | 0.9656% | 3.067% | 0.724% | 1.005% | 1.169% | 1.258% | 2.431% | 0.7368% | 0.8572% | 3.097% | 1.107% | 3.465% | 0.7674% | 2.456%
| PDM3 |  | 0.2295% | -2.032% | -0.7794% | 2.526% | -0.5359% | -0.4663% | -2.554% | 0.5078% | 0.5426% | 2.55% | -0.6158% | 10.46% | 0.7452% | -5.441% | -0.1177% | -8.753%
| Z-Score (Absolute) |  | 0.38 | 1.25 | 0.35 | 2.03 | 0.18 | 0.13 | 1.63 | 0.58 | 0.6 | 2.05 | 0.23 | 7.74 | 0.75 | 3.71 | 0.13 | 6.1
| Analytical results for Ga in OREAS 143 (Certified Value 17.9 ppm) | [17 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 18.7 | 18.16 | 17.2 | 17.8 | 18.8 | 16.8 | 17.7 | 17.65 | 18.4 | 17.4 | 16.7903 | 19 | 18.3202 | 21.4 | 17 | 18.369 | 1
| 1 | 2 | 19 | 19.19 | 17.5 | 17.5 | 18.5 | 17.55 | 16.8 | 18.05 | 18.6 | 17.65 | 16.7017 | 19.1 | 18.0613 | 21.4 | 16 | 18.1029 | 7
| 1 | 3 | 19.2 | 19.06 | 17.7 | 17.8 | 17.6 | 18.3 | 17.7 | 17.7 | 18.1 | 17.4 | 16.5645 | 21 | 18.5935 | 20.8 | 17 | 17.9041 | 16
| 1 | 4 | 19 | 17.96 | 17.31 | 17.8 | 18.1 | 18.25 | 17.9 | 16.35 | 18.5 | 17.5 | 16.3891 | 19.6 | 19.4675 | 20.3 | 16 | 17.9852 | 17.8847
| 1 | 5 | 18.8 | 17.99 | 17.65 | 17.8 | 17.8 | 16.75 | 17.75 | 17.35 | 18.3 | 17.5 | 16.6894 | 19.4 | 18.0028 | 20.9 | 16 | 18.0257 | 90
| 1 | 6 | 19.1 | 18.93 | 17.43 | 17.5 | 18.3 | 17.35 | 17.8 | 17.35 | 17.55 | 18.4 | 16.4313 | 19.6 | 18.5064 | 21.4 | 15 | 18.6965
| Mean |  | 18.9667 | 18.5483 | 17.465 | 17.7 | 18.1833 | 17.5 | 17.6083 | 17.4083 | 18.2417 | 17.6417 | 16.5944 | 19.6167 | 18.492 | 21.0333 | 16.1667 | 18.1806
| Median |  | 19 | 18.545 | 17.465 | 17.8 | 18.2 | 17.45 | 17.725 | 17.5 | 18.35 | 17.5 | 16.6269 | 19.5 | 18.4133 | 21.15 | 16 | 18.0643
| Std Dev. |  | 0.18619 | 0.57059 | 0.19295 | 0.15492 | 0.4446 | 0.67528 | 0.40301 | 0.58001 | 0.38002 | 0.38264 | 0.16036 | 0.72226 | 0.53225 | 0.45019 | 0.75277 | 0.29877
| Rel.Std.Dev. |  | 0.9817% | 3.076% | 1.105% | 0.8753% | 2.445% | 3.859% | 2.289% | 3.332% | 2.083% | 2.169% | 0.9664% | 3.682% | 2.878% | 2.14% | 4.656% | 1.643%
| PDM3 |  | 6.05% | 3.711% | -2.347% | -1.033% | 1.67% | -2.151% | -1.545% | -2.663% | 1.996% | -1.359% | -7.215% | 9.684% | 3.395% | 17.61% | -9.606% | 1.654%
| Z-Score (Absolute) |  | 1.37 | 0.81 | 0.64 | 0.32 | 0.33 | 0.59 | 0.44 | 0.71 | 0.4 | 0.4 | 1.8 | 2.24 | 0.74 | 4.14 | 2.37 | 0.32
| Analytical results for Gd in OREAS 143 (Certified Value 4.63 ppm) | [17 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 4.5 | 4.64 | 4.11 | 4.6 | 5.5598 | 4.7771 | 4.0394 | [9 empty] | 1
| 1 | 2 | 4.5 | 4.71 | 4.19 | 5 | 5.0898 | 4.9056 | 3.9431 | [9 empty] | #N/A
| 1 | 3 | 4.9 | 4.78 | 4.24 | 5 | 5.2651 | 5.0769 | 3.9525 | [9 empty] | 16
| 1 | 4 | 4.4 | 4.63 | 4.05 | 4.8 | 5.0401 | 5.0068 | 3.998 | [9 empty] | 4.626
| 1 | 5 | 4.5 | 4.62 | 4.24 | 4.7 | 5.2716 | 4.7403 | 4.0246 | [9 empty] | 91
| 1 | 6 | 4.7 | 4.83 | 4.19 | 4.6 | 5.2256 | 4.8858 | 4.0617
| Mean |  | 4.5833 | 4.7017 | 4.17 | 4.7833 | 5.242 | 4.8987 | 4.0032
| Median |  | 4.5 | 4.675 | 4.19 | 4.75 | 5.2453 | 4.8957 | 4.0113
| Std Dev. |  | 0.18348 | 0.087502 | 0.075631 | 0.18348 | 0.18255 | 0.12925 | 0.04775
| Rel.Std.Dev. |  | 4.003% | 1.861% | 1.814% | 3.836% | 3.482% | 2.639% | 1.193%
| PDM3 |  | -0.9232% | 1.635% | -9.858% | 3.4% | 13.32% | 5.895% | -13.46%
| Z-Score (Absolute) |  | 0.4 | 0 | 1.82 | 0.28 | 1.85 | 0.67 | 2.39
| Analytical results for Ge in OREAS 143 (Indicative Value 0.15 ppm) | [17 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 14 | 16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 0.19 | 1.6 | < 0.1 | 0.1 | 0.13 | 0.14 | 0.28 | 0.09 | < 0.2 | 1.38 | [6 empty] | 1
| 1 | 2 | 0.2 | 1.6 | < 0.1 | 0.11 | 0.1 | 0.13 | 0.29 | 0.09 | < 0.2 | 1.45 | [6 empty] | 8
| 1 | 3 | 0.19 | 1.6 | < 0.1 | 0.12 | 0.14 | 0.17 | 0.27 | 0.11 | < 0.2 | 1.32 | [6 empty] | 16
| 1 | 4 | 0.18 | 1.6 | < 0.1 | 0.11 | 0.12 | 0.13 | 0.27 | 0.08 | < 0.2 | 1.14 | [6 empty] | 0.1481
| 1 | 5 | 0.19 | 1.7 | < 0.1 | 0.11 | 0.12 | 0.15 | 0.29 | 0.09 | < 0.2 | 1.15 | [6 empty] | 28
| 1 | 6 | 0.18 | 1.6 | 0.1 | 0.12 | 0.14 | 0.13 | 0.25 | 0.11 | < 0.2 | 1.51
| Mean |  | 0.18833 | 1.6167 | 0.1 | 0.11167 | 0.125 | 0.14167 | 0.275 | 0.095 |  | 1.325
| Median |  | 0.19 | 1.6 | 0.1 | 0.11 | 0.125 | 0.135 | 0.275 | 0.09 |  | 1.35
| Std Dev. |  | 0.0075277 | 0.040825 |  | 0.0075277 | 0.015166 | 0.016021 | 0.015166 | 0.012247 |  | 0.15346
| Rel.Std.Dev. |  | 3.997% | 2.525% |  | 6.741% | 12.13% | 11.31% | 5.515% | 12.89% |  | 11.58%
| PDM3 |  | 27.17% | 991.6% | -32.48% | -24.6% | -15.59% | -4.341% | 85.69% | -35.85% |  | 794.7%
| Z-Score (Absolute) |  | 0.79 | 21.43 | 1.08 | 0.31 | 0.12 | 0.12 | 2.05 | 0.55 | 0.48 | 17.22
| Analytical results for Hf in OREAS 143 (Certified Value 2.66 ppm) | [17 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.1 | 3.56 | 5.73 | 2.4 | 2.8 | 2.5 | 2.6 | 2.6 | 2.8 | 2.6 | 2.72 | 2.8849 | 2.8 |  |  |  | 1
| 1 | 2 | 2.5 | 3.33 | 6.09 | 2.5 | 2.8 | 2.7 | 2.7 | 2.6 | 2.8 | 2.6 | 2.82 | 2.9718 | 2.6 |  |  |  | 17
| 1 | 3 | 2.2 | 3.32 | 5.93 | 2.4 | 3 | 2.8 | 2.6 | 2.6 | 3 | 2.5 | 2.9 | 3.0614 | 2.5 |  |  |  | 16
| 1 | 4 | 2.8 | 3.67 | 5.51 | 2.4 | 2.9 | 2.7 | 2.5 | 2.5 | 2.8 | 2.6 | 2.85 | 3.0721 | 2.4 |  |  |  | 2.6627
| 1 | 5 | 2.3 | 3.34 | 5.82 | 2.5 | 2.8 | 2.5 | 2.5 | 2.5 | 2.7 | 2.6 | 2.81 | 2.9192 | 2.5 |  |  |  | 92
| 1 | 6 | 2.4 | 3.47 | 5.99 | 2.5 | 2.8 | 2.5 | 2.8 | 2.6 | 2.9 | 2.8 | 2.74 | 2.988 | 2.6
| Mean |  | 2.3833 | 3.4483 | 5.845 | 2.45 | 2.85 | 2.6167 | 2.6167 | 2.5667 | 2.8333 | 2.6167 | 2.8067 | 2.9829 | 2.5667
| Median |  | 2.35 | 3.405 | 5.875 | 2.45 | 2.8 | 2.6 | 2.6 | 2.6 | 2.8 | 2.6 | 2.815 | 2.9799 | 2.55
| Std Dev. |  | 0.24833 | 0.14442 | 0.20705 | 0.054772 | 0.083666 | 0.13292 | 0.1169 | 0.05164 | 0.10328 | 0.098319 | 0.067429 | 0.074735 | 0.13663
| Rel.Std.Dev. |  | 10.42% | 4.188% | 3.542% | 2.236% | 2.936% | 5.08% | 4.468% | 2.012% | 3.645% | 3.757% | 2.402% | 2.505% | 5.323%
| PDM3 |  | -10.49% | 29.51% | 119.5% | -7.988% | 7.035% | -1.728% | -1.728% | -3.606% | 6.409% | -1.728% | 5.407% | 12.03% | -3.606%
| Z-Score (Absolute) |  | 0.83 | 2.95 | 11.46 | 0.59 | 0.83 | 0 | 0 | 0.18 | 0.77 | 0 | 0.67 | 1.3 | 0.18
| Analytical results for Hg in OREAS 143 (Indicative Value 0.16 ppm) | [17 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 14 | 17 | [14 empty] | ppm
|  |  | 4A*MS | 4A*MS | [14 empty] | 2
2
| 1 | 1 | < 1 | < 0.01 | [14 empty] | 1
| 1 | 2 | < 1 | 0.39 | [14 empty] | 18
| 1 | 3 | < 1 | < 0.01 | [14 empty] | 16
| 1 | 4 | < 1 | 0.29 | [14 empty] | 0.16
| 1 | 5 | < 1 | 0.06 | [14 empty] | 26
| 1 | 6 | < 1 | 0.21
| Mean |  |  | 0.2375
| Median |  |  | 0.25
| Std Dev. |  |  | 0.13937
| Rel.Std.Dev. |  |  | 58.68%
| PDM3 |  |  | 48.44%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 143 (Certified Value 0.52 ppm) | [17 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 0.5 | 0.58 | 0.48 | 0.6 | 0.47098 | 0.5285 | 0.4528 | [9 empty] | 1
| 1 | 2 | 0.5 | 0.58 | 0.46 | 0.6 | 0.47637 | 0.55963 | 0.4446 | [9 empty] | #N/A
| 1 | 3 | 0.6 | 0.6 | 0.47 | 0.6 | 0.47632 | 0.55672 | 0.4613 | [9 empty] | 16
| 1 | 4 | 0.4 | 0.56 | 0.44 | 0.7 | 0.48901 | 0.56415 | 0.4505 | [9 empty] | 0.52131
| 1 | 5 | 0.5 | 0.6 | 0.47 | 0.6 | 0.4377 | 0.54381 | 0.4737 | [9 empty] | 93
| 1 | 6 | 0.5 | 0.58 | 0.48 | 0.6 | 0.45268 | 0.56525 | 0.4912
| Mean |  | 0.5 | 0.58333 | 0.46667 | 0.61667 | 0.46718 | 0.55301 | 0.46235
| Median |  | 0.5 | 0.58 | 0.47 | 0.6 | 0.47365 | 0.55818 | 0.45705
| Std Dev. |  | 0.063246 | 0.015055 | 0.015055 | 0.040825 | 0.018646 | 0.014262 | 0.017372
| Rel.Std.Dev. |  | 12.65% | 2.581% | 3.226% | 6.62% | 3.991% | 2.579% | 3.757%
| PDM3 |  | -4.089% | 11.9% | -10.48% | 18.29% | -10.38% | 6.08% | -11.31%
| Z-Score (Absolute) |  | 0 | 1.49 | 0.6 | 2.09 | 0.59 | 0.95 | 0.67
| Analytical results for In in OREAS 143 (Certified Value 0.87 ppm) | [17 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 16 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 1.15 | 1.02 | 1.01 | 0.8 | 0.808 | 0.866 | 0.607 | 0.739 | 1.02 | 0.66331 | 0.62 | 1.1755 | 0.904 |  |  |  | 1
| 1 | 2 | 0.81 | 1.01 | 0.93 | 0.8 | 0.834 | 0.849 | 0.934 | 2.55 | 1.21 | 0.63339 | 0.62 | 1.2057 | 1.2 |  |  |  | 2
| 1 | 3 | 2.19 | 0.99 | 0.98 | 0.8 | 0.83 | 0.857 | 0.776 | 0.808 | 0.879 | 0.63388 | 0.54 | 1.1232 | 1.11 |  |  |  | 16
| 1 | 4 | 1.97 | 0.93 | 0.92 | 0.8 | 0.955 | 0.871 | 0.91 | 0.92 | 0.547 | 0.62529 | 0.68 | 1.0424 | 1.43 |  |  |  | 0.87151
| 1 | 5 | 0.84 | 0.97 | 0.96 | 0.8 | 0.803 | 0.859 | 0.61 | < 0.005 | 1.045 | 0.60802 | 0.44 | 1.0978 | 1.01 |  |  |  | 94
| 1 | 6 | 0.71 | 1.01 | 0.92 | 0.9 | 0.843 | 0.829 | 0.108 | < 0.005 | 0.743 | 0.64556 | 0.65 | 1.1963 | 1.27
| Mean |  | 1.2783 | 0.98833 | 0.95333 | 0.81667 | 0.8455 | 0.85517 | 0.6575 | 1.2542 | 0.90733 | 0.63491 | 0.59167 | 1.1402 | 1.154
| Median |  | 0.995 | 1 | 0.945 | 0.8 | 0.832 | 0.858 | 0.693 | 0.864 | 0.9495 | 0.63363 | 0.62 | 1.1494 | 1.155
| Std Dev. |  | 0.64195 | 0.033714 | 0.036697 | 0.040825 | 0.055824 | 0.014892 | 0.30365 | 0.86705 | 0.23702 | 0.018651 | 0.087731 | 0.063775 | 0.18811
| Rel.Std.Dev. |  | 50.22% | 3.411% | 3.849% | 4.999% | 6.602% | 1.741% | 46.18% | 69.13% | 26.12% | 2.938% | 14.83% | 5.594% | 16.3%
| PDM3 |  | 46.68% | 13.41% | 9.389% | -6.292% | -2.984% | -1.875% | -24.56% | 43.92% | 4.111% | -27.15% | -32.11% | 30.83% | 32.41%
| Z-Score (Absolute) |  | 1.62 | 0.44 | 0.29 | NA | 0.15 | 0.11 | 0.91 | 0.18 | 0.11 | 1 | 1.18 | 1.06 | 1.11
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 143 (Certified Value 2.36 wt.%) | [17 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 2.31 | 2.29 | 2.32 | 2.16 | 2.3 | 2.32 | 2.29 | 2.34 | 2.38 | 2.395 | 2.438 | 2.43 | 2.3385 | 2.3 | 2.1376 |  | 1
| 1 | 2 | 2.26 | 2.36 | 2.3 | 2.28 | 2.32 | 2.32 | 2.35 | 2.37 | 2.37 | 2.423 | 2.4379 | 2.44 | 2.3392 | 2.28 | 2.1721 |  | #N/A
| 1 | 3 | 2.32 | 2.34 | 2.4 | 2 | 2.32 | 2.31 | 2.32 | 2.37 | 2.33 | 2.419 | 2.4463 | 2.6 | 2.3789 | 2.59 | 2.085 |  | 16
| 1 | 4 | 2.32 | 2.35 | 2.39 | 2.29 | 2.33 | 2.3 | 2.31 | 2.4 | 2.33 | 2.44 | 2.4237 | 2.48 | 2.4063 | 2.56 | 2.2259 |  | 2.3554
| 1 | 5 | 2.31 | 2.34 | 2.31 | 2.25 | 2.3 | 2.35 | 2.29 | 2.32 | 2.38 | 2.387 | 2.4534 | 2.4 | 2.3586 | 2.39 | 2.2106 |  | 95
| 1 | 6 | 2.31 | 2.35 | 2.35 | 2.26 | 2.36 | 2.34 | 2.29 | 2.36 | 2.49 | 2.408 | 2.4222 | 2.44 | 2.3775 | 2.54 | 2.0073
| Mean |  | 2.305 | 2.3383 | 2.345 | 2.2067 | 2.3217 | 2.3233 | 2.3083 | 2.36 | 2.38 | 2.412 | 2.4369 | 2.465 | 2.3665 | 2.4433 | 2.1398
| Median |  | 2.31 | 2.345 | 2.335 | 2.255 | 2.32 | 2.32 | 2.3 | 2.365 | 2.375 | 2.4135 | 2.4379 | 2.44 | 2.368 | 2.465 | 2.1549
| Std Dev. |  | 0.022583 | 0.024833 | 0.042308 | 0.1113 | 0.022286 | 0.018619 | 0.024014 | 0.027568 | 0.058652 | 0.019411 | 0.012259 | 0.070922 | 0.026275 | 0.1375 | 0.082497
| Rel.Std.Dev. |  | 0.9797% | 1.062% | 1.804% | 5.044% | 0.9599% | 0.8014% | 1.04% | 1.168% | 2.464% | 0.8048% | 0.503% | 2.877% | 1.11% | 5.628% | 3.855%
| PDM3 |  | -2.138% | -0.723% | -0.4399% | -6.313% | -1.431% | -1.36% | -1.997% | 0.1969% | 1.046% | 2.405% | 3.462% | 4.655% | 0.4722% | 3.735% | -9.153%
| Z-Score (Absolute) |  | 0.74 | 0.12 | 0 | 2.54 | 0.43 | 0.4 | 0.67 | 0.28 | 0.64 | 1.23 | 1.69 | 2.21 | 0.4 | 1.81 | 3.77
| Analytical results for La in OREAS 143 (Certified Value 33.2 ppm) | [17 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 28.8 | 29.95 | 30.39 | 34.4 | 33 | 31.5 | 36.2 | 34.6 | 35.3 | 32.78 | 31.9595 | 44.4 | 33.7446 | 20.2 | 34.0333 |  | 1
| 1 | 2 | 32.1 | 30.45 | 30.66 | 35.9 | 33.8 | 31.9 | 37.1 | 35.5 | 35.3 | 32.81 | 32.2724 | 43.9 | 34.0553 | 19.7 | 37.1192 |  | 3
| 1 | 3 | 33.6 | 30.65 | 31.66 | 35.3 | 35.3 | 30.7 | 36.6 | 34.8 | 35 | 32.985 | 32.362 | 46.2 | 34.6797 | 20.3 | 33.0743 |  | 16
| 1 | 4 | 30 | 29.87 | 29.3 | 35.9 | 35.5 | 31.2 | 34.3 | 35.2 | 34.6 | 32.787 | 32.1649 | 43.7 | 35.1222 | 20.3 | 33.5187 |  | 33.2382
| 1 | 5 | 28.9 | 30.87 | 31.53 | 34 | 32.1 | 33.1 | 34.9 | 35.3 | 35.1 | 32.496 | 31.9477 | 44 | 33.9147 | 20.1 | 33.1373 |  | 96
| 1 | 6 | 30.9 | 30.93 | 30.84 | 35.1 | 33.5 | 32.5 | 35.2 | 34.2 | 36.5 | 32.858 | 31.9016 | 44.1 | 34.4215 | 20.3 | 34.6321
| Mean |  | 30.7167 | 30.4533 | 30.73 | 35.1 | 33.8667 | 31.8167 | 35.7167 | 34.9333 | 35.3 | 32.786 | 32.1013 | 44.3833 | 34.323 | 20.15 | 34.2525
| Median |  | 30.45 | 30.55 | 30.75 | 35.2 | 33.65 | 31.7 | 35.7 | 35 | 35.2 | 32.7985 | 32.0622 | 44.05 | 34.2384 | 20.25 | 33.776
| Std Dev. |  | 1.8841 | 0.45456 | 0.85818 | 0.77717 | 1.3216 | 0.87731 | 1.0834 | 0.48854 | 0.64187 | 0.16093 | 0.19228 | 0.9196 | 0.51931 | 0.23452 | 1.5219
| Rel.Std.Dev. |  | 6.134% | 1.493% | 2.793% | 2.214% | 3.902% | 2.757% | 3.033% | 1.398% | 1.818% | 0.4909% | 0.599% | 2.072% | 1.513% | 1.164% | 4.443%
| PDM3 |  | -7.586% | -8.378% | -7.546% | 5.601% | 1.891% | -4.277% | 7.457% | 5.1% | 6.203% | -1.36% | -3.42% | 33.53% | 3.264% | -39.38% | 3.052%
| Z-Score (Absolute) |  | 1.2 | 1.3 | 1.2 | 0.47 | 0 | 0.78 | 0.71 | 0.41 | 0.55 | 0.41 | 0.67 | 4.02 | 0.17 | 5.24 | 0.15
| Analytical results for Li in OREAS 143 (Certified Value 59 ppm) | [17 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  | 0
1
| 1 | 1 | 61 | 56.43 | 56.7 | 61 | 60.1 | 60.1 | 59.7 | 61.6 | 60.6 | 63.74 | 53.485 | 55 | 57.5298 | 41.4 | 44.83 |  | 1
| 1 | 2 | 55 | 55.95 | 54.5 | 62 | 60 | 60.1 | 61.5 | 62.8 | 61.6 | 63.797 | 53.2268 | 52 | 57.4928 | 43.1 | 45.375 |  | #N/A
| 1 | 3 | 60 | 56.14 | 56.5 | 62 | 60.4 | 59.7 | 60.6 | 62.3 | 58.9 | 64.415 | 55.2095 | 58 | 58.2387 | 39.8 | 42.68 |  | 16
| 1 | 4 | 59 | 53 | 55.5 | 61 | 60.6 | 59.4 | 59.7 | 63.6 | 59 | 64.581 | 55.8279 | 54 | 58.5713 | 37.8 | 47.13 |  | 58.9647
| 1 | 5 | 58 | 56.2 | 55.5 | 61 | 60.2 | 60.9 | 59.9 | 61.1 | 60.2 | 63.188 | 54.7675 | 53 | 58.2905 | 39.5 | 44.095 |  | 97
| 1 | 6 | 59 | 54.41 | 56.9 | 61 | 61.2 | 60.4 | 59.9 | 62.2 | 63.4 | 64.04 | 55.7676 | 55 | 57.8203 | 43.2 | 42.485
| Mean |  | 58.6667 | 55.355 | 55.9333 | 61.3333 | 60.4167 | 60.1 | 60.2167 | 62.2667 | 60.6167 | 63.9602 | 54.714 | 54.5 | 57.9905 | 40.8 | 44.4325
| Median |  | 59 | 56.045 | 56 | 61 | 60.3 | 60.1 | 59.9 | 62.25 | 60.4 | 63.9185 | 54.9885 | 54.5 | 58.0295 | 40.6 | 44.4625
| Std Dev. |  | 2.0656 | 1.3623 | 0.92448 | 0.5164 | 0.44008 | 0.52536 | 0.7111 | 0.88015 | 1.6999 | 0.50399 | 1.1244 | 2.0736 | 0.44223 | 2.1494 | 1.7493
| Rel.Std.Dev. |  | 3.521% | 2.461% | 1.653% | 0.842% | 0.7284% | 0.8741% | 1.181% | 1.414% | 2.804% | 0.788% | 2.055% | 3.805% | 0.7626% | 5.268% | 3.937%
| PDM3 |  | -0.5054% | -6.122% | -5.141% | 4.017% | 2.462% | 1.925% | 2.123% | 5.6% | 2.802% | 8.472% | -7.209% | -7.572% | -1.652% | -30.81% | -24.65%
| Z-Score (Absolute) |  | 0 | 0.82 | 0.67 | 0.66 | 0.43 | 0.35 | 0.38 | 0.89 | 0.48 | 1.31 | 0.98 | 1.03 | 0.17 | 4.41 | 3.51
| Analytical results for Lu in OREAS 143 (Certified Value 0.19 ppm) | [17 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 13 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 0.17 | < 0.5 | 0.17 | 0.3 | 0.2 | 0.21107 | 0.1735 | [9 empty] | 1
| 1 | 2 | 0.16 | < 0.5 | 0.19 | 0.2 | 0.2 | 0.19976 | 0.1705 | [9 empty] | 13
| 1 | 3 | 0.22 | < 0.5 | 0.18 | 0.3 | 0.22 | 0.20604 | 0.1764 | [9 empty] | 16
| 1 | 4 | 0.17 | < 0.5 | 0.18 | 0.3 | 0.21 | 0.20649 | 0.1747 | [9 empty] | 0.19187
| 1 | 5 | 0.22 | < 0.5 | 0.18 | 0.3 | 0.21 | 0.20223 | 0.1808 | [9 empty] | 98
| 1 | 6 | 0.19 | < 0.5 | 0.18 | 0.3 | 0.2 | 0.20812 | 0.1964
| Mean |  | 0.18833 |  | 0.18 | 0.28333 | 0.20667 | 0.20562 | 0.17872
| Median |  | 0.18 |  | 0.18 | 0.3 | 0.205 | 0.20627 | 0.17555
| Std Dev. |  | 0.026394 |  | 0.0063246 | 0.040825 | 0.008165 | 0.0040688 | 0.0093088
| Rel.Std.Dev. |  | 14.01% |  | 3.514% | 14.41% | 3.951% | 1.979% | 5.209%
| PDM3 |  | -1.842% |  | -6.185% | 47.67% | 7.713% | 7.168% | -6.854%
| Z-Score (Absolute) |  | 0.45 | 1.17 | 0.67 | 2.05 | 0.03 | 0 | 0.71
| Analytical results for Mg in OREAS 143 (Certified Value 1.15 wt.%) | [17 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 1.14 | 1.16 | 1.14 | 1.15 | 1.14 | 1.14 | 1.12 | 1.17 | 1.19 | 1.144 | 1.1542 | 1.11 | 1.1434 | 1.2 | 1.1185 | 1.1368 | 1
| 1 | 2 | 1.13 | 1.18 | 1.13 | 1.18 | 1.13 | 1.13 | 1.16 | 1.18 | 1.2 | 1.141 | 1.1445 | 1.12 | 1.1654 | 1.17 | 1.1141 | 1.1482 | #N/A
| 1 | 3 | 1.16 | 1.15 | 1.17 | 1.15 | 1.14 | 1.13 | 1.14 | 1.18 | 1.16 | 1.146 | 1.1568 | 1.19 | 1.171 | 1.22 | 1.1228 | 1.1028 | 16
| 1 | 4 | 1.14 | 1.13 | 1.18 | 1.25 | 1.14 | 1.12 | 1.14 | 1.21 | 1.16 | 1.147 | 1.1274 | 1.14 | 1.1812 | 1.16 | 1.1374 | 1.1756 | 1.1509
| 1 | 5 | 1.14 | 1.17 | 1.13 | 1.21 | 1.14 | 1.15 | 1.13 | 1.16 | 1.18 | 1.134 | 1.1561 | 1.11 | 1.171 | 1.13 | 1.1145 | 1.1284 | 99
| 1 | 6 | 1.15 | 1.16 | 1.17 | 1.19 | 1.15 | 1.13 | 1.14 | 1.18 | 1.23 | 1.147 | 1.1333 | 1.12 | 1.1737 | 1.12 | 1.114 | 1.0867
| Mean |  | 1.1433 | 1.1583 | 1.1533 | 1.1883 | 1.14 | 1.1333 | 1.1383 | 1.18 | 1.1867 | 1.1432 | 1.1454 | 1.1317 | 1.1676 | 1.1667 | 1.1202 | 1.1297
| Median |  | 1.14 | 1.16 | 1.155 | 1.185 | 1.14 | 1.13 | 1.14 | 1.18 | 1.185 | 1.145 | 1.1494 | 1.12 | 1.171 | 1.165 | 1.1165 | 1.1326
| Std Dev. |  | 0.010328 | 0.017224 | 0.022509 | 0.038166 | 0.0063246 | 0.010328 | 0.013292 | 0.016733 | 0.026583 | 0.0050365 | 0.012589 | 0.030605 | 0.012945 | 0.038816 | 0.0090954 | 0.031848
| Rel.Std.Dev. |  | 0.9033% | 1.487% | 1.952% | 3.212% | 0.5548% | 0.9113% | 1.168% | 1.418% | 2.24% | 0.4406% | 1.099% | 2.704% | 1.109% | 3.327% | 0.8119% | 2.819%
| PDM3 |  | -0.6541% | 0.6493% | 0.2148% | 3.256% | -0.9437% | -1.523% | -1.089% | 2.532% | 3.111% | -0.6686% | -0.4765% | -1.668% | 1.456% | 1.373% | -2.663% | -1.835%
| Z-Score (Absolute) |  | 0.05 | 0.71 | 0.45 | 2.22 | 0.22 | 0.56 | 0.3 | 1.8 | 2.14 | 0.06 | 0.05 | 0.64 | 1.18 | 1.13 | 1.22 | 0.74
| Analytical results for Mn in OREAS 143 (Certified Value 0.045 wt.%) | [17 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 3
3
| 1 | 1 | 0.0455 | 0.04544 | 0.0453 | 0.0469 | 0.0438 | 0.0454 | 0.0434 | 0.0467 | 0.0445 | 0.044758 | 0.040652 | 0.0457 | 0.043825 | 0.0506 | 0.0414 | 0.04132 | 1
| 1 | 2 | 0.0446 | 0.044848 | 0.0447 | 0.0438 | 0.0435 | 0.0438 | 0.0445 | 0.0469 | 0.0466 | 0.043992 | 0.040587 | 0.0457 | 0.044242 | 0.0494 | 0.0411 | 0.044154 | #N/A
| 1 | 3 | 0.0462 | 0.044034 | 0.0455 | 0.0474 | 0.0443 | 0.0444 | 0.0437 | 0.0462 | 0.0436 | 0.044242 | 0.040614 | 0.049 | 0.045134 | 0.0485 | 0.0406 | 0.038545 | 16
| 1 | 4 | 0.0457 | 0.04313 | 0.0456 | 0.0457 | 0.044 | 0.0438 | 0.0452 | 0.0485 | 0.0431 | 0.044333 | 0.040724 | 0.0464 | 0.04559 | 0.0477 | 0.0405 | 0.046278 | 0.044617
| 1 | 5 | 0.0458 | 0.043914 | 0.0454 | 0.0472 | 0.0441 | 0.0454 | 0.0439 | 0.0461 | 0.0442 | 0.043859 | 0.040831 | 0.0451 | 0.044285 | 0.048 | 0.0416 | 0.043528 | 100
| 1 | 6 | 0.0459 | 0.0443 | 0.0453 | 0.0462 | 0.0445 | 0.0439 | 0.0444 | 0.0467 | 0.0456 | 0.0446 | 0.040577 | 0.0455 | 0.045009 | 0.0507 | 0.0414 | 0.043182
| Mean |  | 0.045617 | 0.044278 | 0.0453 | 0.0462 | 0.044033 | 0.04445 | 0.044183 | 0.04685 | 0.0446 | 0.044297 | 0.040664 | 0.046233 | 0.044681 | 0.04915 | 0.0411 | 0.042835
| Median |  | 0.04575 | 0.044167 | 0.04535 | 0.04655 | 0.04405 | 0.04415 | 0.04415 | 0.0467 | 0.04435 | 0.044287 | 0.040633 | 0.0457 | 0.044647 | 0.04895 | 0.04125 | 0.043355
| Std Dev. |  | 0.00054924 | 0.00079819 | 0.00031623 | 0.0013372 | 0.0003559 | 0.00076877 | 0.00064936 | 0.0008666 | 0.0012977 | 0.00034468 | 9.7602e-05 | 0.0014194 | 0.0006666 | 0.0012973 | 0.00045607 | 0.0026419
| Rel.Std.Dev. |  | 1.204% | 1.803% | 0.6981% | 2.894% | 0.8083% | 1.73% | 1.47% | 1.85% | 2.91% | 0.7781% | 0.24% | 3.07% | 1.492% | 2.639% | 1.11% | 6.168%
| PDM3 |  | 2.242% | -0.7595% | 1.532% | 3.549% | -1.307% | -0.3733% | -0.971% | 5.006% | -0.03708% | -0.7152% | -8.859% | 3.624% | 0.1438% | 10.16% | -7.882% | -3.994%
| Z-Score (Absolute) |  | 0.79 | 0.18 | 0.56 | 1.21 | 0.36 | 0.05 | 0.25 | 1.68 | 0.05 | 0.16 | 2.79 | 1.23 | 0.11 | 3.34 | 2.47 | 1.22
| Analytical results for Mo in OREAS 143 (Certified Value 1.32 ppm) | [17 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 2 | 1.53 | 1.3 | < 1 | 1.32 | 1.28 | 1.3 | 1.43 | 1.28 | 1.301 | 1.46 | 2.7329 | 1.31 | 1.5 |  |  | 1
| 1 | 2 | 1 | 1.59 | 1.3 | 2 | 1.28 | 1.31 | 1.38 | 1.34 | 1.27 | < 1 | 1.56 | 2.1967 | 1.23 | 1.5 |  |  | 14
| 1 | 3 | 2 | 1.63 | 1.3 | 3 | 1.3 | 5.28 | 1.42 | 1.34 | 1.22 | 1.111 | 1.64 | 2.1134 | 1.32 | 1.3 |  |  | 16
| 1 | 4 | 1 | 1.63 | 1.2 | < 1 | 1.54 | 1.31 | 1.3 | 1.32 | 1.24 | < 1 | 1.56 | 2.4696 | 1.13 | 1.3 | 2.37 |  | 1.3203
| 1 | 5 | 2 | 1.61 | 1.3 | 1 | 1.23 | 1.38 | 1.2 | 4.06 | 1.36 | < 1 | 1.43 | 1.9436 | 1.24 | 1.3 | 2.55 |  | 101
| 1 | 6 | 1 | 1.53 | 1.3 | 1 | 1.38 | 1.3 | 1.44 | 1.32 | 1.32 | < 1 | 1.42 | 2.0601 | 1.3 | 1.4 | 2.85
| Mean |  | 1.5 | 1.5867 | 1.2833 | 1.75 | 1.3417 | 1.9767 | 1.34 | 1.8017 | 1.2817 | 1.206 | 1.5117 | 2.2527 | 1.255 | 1.3833 | 2.59
| Median |  | 1.5 | 1.6 | 1.3 | 1.5 | 1.31 | 1.31 | 1.34 | 1.34 | 1.275 | 1.206 | 1.51 | 2.155 | 1.27 | 1.35 | 2.55
| Std Dev. |  | 0.54772 | 0.046332 | 0.040825 | 0.95743 | 0.10889 | 1.6186 | 0.090333 | 1.1071 | 0.051543 | 0.13435 | 0.088185 | 0.29422 | 0.071764 | 0.098319 | 0.24249
| Rel.Std.Dev. |  | 36.51% | 2.92% | 3.181% | 54.71% | 8.116% | 81.89% | 6.741% | 61.45% | 4.022% | 11.14% | 5.834% | 13.06% | 5.718% | 7.107% | 9.362%
| PDM3 |  | 13.61% | 20.17% | -2.802% | 32.54% | 1.616% | 49.71% | 1.49% | 36.46% | -2.929% | -8.659% | 14.49% | 70.62% | -4.948% | 4.772% | 96.16%
| Z-Score (Absolute) |  | NA | 1.07 | 0.53 | NA | 0.22 | 3.12 | 0.23 | 2.2 | 0.53 | 3.4 | 0.67 | 4.57 | 0.67 | 0 | 6.34
Results from laboratories 1 and 5 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 143 (Certified Value 0.926 wt.%) | [17 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.91 | 0.89 | 0.901 | 0.86 | 0.94 | 0.91 | 0.94 | 0.93 | 0.93 | 0.95 | 0.97 | 0.92077 | 0.86 | 0.8781 | 0.9631 |  | 1
| 1 | 2 | 0.9 | 0.92 | 0.893 | 0.88 | 0.94 | 0.91 | 0.97 | 0.94 | 0.93 | 0.961 | 0.98 | 0.92532 | 0.84 | 0.8708 | 0.97594 |  | 15
| 1 | 3 | 0.92 | 0.92 | 0.918 | 0.83 | 0.93 | 0.91 | 0.96 | 0.94 | 0.91 | 0.961 | 1.04 | 0.94009 | 0.82 | 0.8816 | 0.9345 |  | 16
| 1 | 4 | 0.91 | 0.91 | 0.903 | 0.95 | 0.95 | 0.9 | 0.92 | 0.95 | 0.91 | 0.966 | 0.99 | 0.94355 | 0.8 | 0.8891 | 1.0117 |  | 0.92567
| 1 | 5 | 0.9 | 0.91 | 0.9 | 0.84 | 0.93 | 0.92 | 0.95 | 0.92 | 0.93 | 0.95 | 0.96 | 0.93204 | 0.81 | 0.8747 | 0.95747 |  | 102
| 1 | 6 | 0.9 | 0.92 | 0.914 | 0.83 | 0.95 | 0.91 | 0.95 | 0.94 | 0.98 | 0.957 | 0.97 | 0.94056 | 0.84 | 0.8696 | 0.92918
| Mean |  | 0.90667 | 0.91167 | 0.90483 | 0.865 | 0.94 | 0.91 | 0.94833 | 0.93667 | 0.93167 | 0.9575 | 0.985 | 0.93372 | 0.82833 | 0.87732 | 0.96198
| Median |  | 0.905 | 0.915 | 0.902 | 0.85 | 0.94 | 0.91 | 0.95 | 0.94 | 0.93 | 0.959 | 0.975 | 0.93607 | 0.83 | 0.8764 | 0.96028
| Std Dev. |  | 0.008165 | 0.01169 | 0.0093684 | 0.045935 | 0.0089443 | 0.0063246 | 0.017224 | 0.010328 | 0.025626 | 0.006473 | 0.02881 | 0.0092185 | 0.022286 | 0.007303 | 0.030058
| Rel.Std.Dev. |  | 0.9005% | 1.282% | 1.035% | 5.31% | 0.9515% | 0.695% | 1.816% | 1.103% | 2.751% | 0.676% | 2.925% | 0.9873% | 2.69% | 0.8324% | 3.125%
| PDM3 |  | -2.053% | -1.513% | -2.251% | -6.554% | 1.548% | -1.693% | 2.449% | 1.188% | 0.6481% | 3.439% | 6.41% | 0.8701% | -10.52% | -5.223% | 3.922%
| Z-Score (Absolute) |  | 0.67 | 0.54 | 0.72 | 1.8 | 0.22 | 0.58 | 0.45 | 0.13 | 0 | 0.7 | 1.44 | 0.06 | 2.79 | 1.47 | 0.82
| Analytical results for Nb in OREAS 143 (Certified Value 11.3 ppm) | [17 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  |  | 1
2
| 1 | 1 | 12.1 | 11.72 | 7.36 | 9.3 | 11 | 10.5 | 11 | 12.4 | 12.8 | 11.5 | 10.7 | 10.7733 | 11.7 | 12.14 |  |  | 1
| 1 | 2 | 10.1 | 12.33 | 7.05 | 9.3 | 12.9 | 10.8 | 11 | 12.1 | 13 | 11.4 | 10.4 | 11.0787 | 11.7 | 10.1 |  |  | 4
| 1 | 3 | 9.4 | 12.18 | 7.56 | 9.3 | 11.4 | 11.2 | 11 | 12.1 | 12.7 | 11.2 | 11.7 | 11.2067 | 11.4 | 11.93 |  |  | 16
| 1 | 4 | 9.7 | 12.26 | 7.48 | 9.5 | 10.8 | 11.3 | 10.7 | 11.1 | 13 | 11.2 | 10.8 | 11.1618 | 11 | 13.89 |  |  | 11.2766
| 1 | 5 | 9.5 | 11.79 | 7.33 | 9.3 | 10.6 | 10.2 | 11.4 | 11.8 | 13.1 | 11.2 | 11 | 10.9124 | 10.9 | 14.99 |  |  | 103
| 1 | 6 | 9.8 | 11.84 | 7.65 | 9.6 | 13 | 10.8 | 11.6 | 12.4 | 12.6 | 11.9 | 10.5 | 11.1285 | 11.7 | 63.56
| Mean |  | 10.1 | 12.02 | 7.405 | 9.3833 | 11.6167 | 10.8 | 11.1167 | 11.9833 | 12.8667 | 11.4 | 10.85 | 11.0436 | 11.4 | 21.1017
| Median |  | 9.75 | 12.01 | 7.42 | 9.3 | 11.2 | 10.8 | 11 | 12.1 | 12.9 | 11.3 | 10.75 | 11.1036 | 11.55 | 13.015
| Std Dev. |  | 1.01 | 0.26631 | 0.21135 | 0.13292 | 1.0666 | 0.41473 | 0.32506 | 0.48751 | 0.19664 | 0.27568 | 0.46797 | 0.1669 | 0.36878 | 20.8688
| Rel.Std.Dev. |  | 10% | 2.216% | 2.854% | 1.417% | 9.182% | 3.84% | 2.924% | 4.068% | 1.528% | 2.418% | 4.313% | 1.511% | 3.235% | 98.9%
| PDM3 |  | -10.43% | 6.593% | -34.33% | -16.79% | 3.016% | -4.226% | -1.418% | 6.268% | 14.1% | 1.095% | -3.783% | -2.066% | 1.095% | 87.13%
| Z-Score (Absolute) |  | 1.32 | 0.87 | 4.39 | 2.14 | 0.41 | 0.52 | 0.16 | 0.83 | 1.83 | 0.16 | 0.47 | 0.24 | 0.16 | 11.22
| Analytical results for Nd in OREAS 143 (Certified Value 28.3 ppm) | [17 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 14 | 17 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 28.1 | 28.5 | 27.29 | 29 | 29.9476 | 25 | 27.5097 | [9 empty] | 1
| 1 | 2 | 30.7 | 28.02 | 27.89 | 30.6 | 31.1657 | 23 | 27.8176 | [9 empty] | #N/A
| 1 | 3 | 32.5 | 26.85 | 28.16 | 31 | 31.3682 | 21 | 26.9764 | [9 empty] | 16
| 1 | 4 | 30 | 26.98 | 26.39 | 30.9 | 31.5376 | 26 | 27.486 | [9 empty] | 28.28
| 1 | 5 | 27.8 | 29.08 | 28.45 | 29.5 | 30.5503 | 23 | 28.321 | [9 empty] | 104
| 1 | 6 | 29.7 | 29.44 | 27.27 | 29.9 | 31.0727 | 23 | 28.9871
| Mean |  | 29.8 | 28.145 | 27.575 | 30.15 | 30.9404 | 23.5 | 27.8496
| Median |  | 29.85 | 28.26 | 27.59 | 30.25 | 31.1192 | 23 | 27.6636
| Std Dev. |  | 1.7344 | 1.0702 | 0.7464 | 0.81179 | 0.59088 | 1.7607 | 0.71097
| Rel.Std.Dev. |  | 5.82% | 3.803% | 2.707% | 2.692% | 1.91% | 7.492% | 2.553%
| PDM3 |  | 5.375% | -0.4774% | -2.493% | 6.612% | 9.407% | -16.9% | -1.522%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.23 | 0.82 | 1.14 | 1.89 | 0.12
| Analytical results for Ni in OREAS 143 (Certified Value 33.9 ppm) | [17 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 44 | 33.52 | 34 | 34 | 30.4 | 33.2 | 32.4 | 35.6 | 32.4 | 36.423 | 31.7235 | 38.9 | 34.231 | 34.6 | 30.2 | 32.69 | 1
| 1 | 2 | 41 | 34.84 | 33 | 34 | 31.6 | 31.8 | 32.9 | 35.7 | 32.8 | 36.456 | 31.642 | 38.4 | 34.3312 | 37.1 | 31.1 | 30.86 | #N/A
| 1 | 3 | 34 | 33.62 | 34 | 33 | 32.8 | 35.9 | 32.6 | 36.1 | 32.3 | 36.094 | 31.4135 | 41.9 | 34.5119 | 35.2 | 31 | 29.64 | 16
| 1 | 4 | 36 | 33.52 | 34 | 34 | 33.1 | 33.5 | 32.7 | 35.2 | 32.2 | 36.983 | 31.3665 | 39.1 | 34.5259 | 36.1 | 31.9 | 30.99 | 33.8525
| 1 | 5 | 35 | 33.47 | 34 | 33 | 29.9 | 34.6 | 31.6 | 57.1 | 32.4 | 35.543 | 31.7785 | 37.4 | 35.2366 | 34.7 | 30.6 | 30.54 | 105
| 1 | 6 | 38 | 32.56 | 34 | 33 | 31.9 | 33.1 | 32.2 | 34.1 | 33.7 | 36.509 | 31.3585 | 39.3 | 35.3142 | 36.5 | 30.5 | 28.09
| Mean |  | 38 | 33.5883 | 33.8333 | 33.5 | 31.6167 | 33.6833 | 32.4 | 38.9667 | 32.6333 | 36.3347 | 31.5471 | 39.1667 | 34.6918 | 35.7 | 30.8833 | 30.4683
| Median |  | 37 | 33.52 | 34 | 33.5 | 31.75 | 33.35 | 32.5 | 35.65 | 32.4 | 36.4395 | 31.5278 | 39 | 34.5189 | 35.65 | 30.8 | 30.7
| Std Dev. |  | 3.8471 | 0.7278 | 0.40825 | 0.54772 | 1.2734 | 1.4077 | 0.46043 | 8.9097 | 0.56095 | 0.48123 | 0.18958 | 1.5016 | 0.46613 | 1.0218 | 0.59805 | 1.5303
| Rel.Std.Dev. |  | 10.12% | 2.167% | 1.207% | 1.635% | 4.028% | 4.179% | 1.421% | 22.86% | 1.719% | 1.324% | 0.6009% | 3.834% | 1.344% | 2.862% | 1.936% | 5.023%
| PDM3 |  | 12.25% | -0.7804% | -0.05666% | -1.041% | -6.605% | -0.4998% | -4.291% | 15.11% | -3.601% | 7.332% | -6.81% | 15.7% | 2.479% | 5.457% | -8.771% | -9.997%
| Z-Score (Absolute) |  | 1.44 | 0.02 | 0.07 | 0.04 | 0.67 | 0.02 | 0.41 | 1.76 | 0.33 | 0.89 | 0.69 | 1.83 | 0.35 | 0.68 | 0.91 | 1.05
| Analytical results for P in OREAS 143 (Certified Value 0.058 wt.%) | [17 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 3
3
| 1 | 1 | 0.06 | 0.0587 | 0.057 | 0.054 | 0.057 | 0.058 | 0.056 | 0.058 | 0.058 | 0.058 | 0.07856 | 0.06 | 0.05519 | 0.061 | 0.0597 | 0.049033 | 1
| 1 | 2 | 0.06 | 0.0577 | 0.056 | 0.056 | 0.056 | 0.058 | 0.058 | 0.06 | 0.057 | 0.057 | 0.078 | 0.06 | 0.057256 | 0.059 | 0.0582 | 0.051286 | #N/A
| 1 | 3 | 0.06 | 0.0587 | 0.058 | 0.054 | 0.057 | 0.058 | 0.057 | 0.059 | 0.056 | 0.06 | 0.07521 | 0.06 | 0.055354 | 0.061 | 0.0563 | 0.049641 | 16
| 1 | 4 | 0.06 | 0.0571 | 0.058 | 0.054 | 0.057 | 0.057 | 0.058 | 0.06 | 0.057 | 0.059 | 0.07818 | 0.06 | 0.05683 | 0.059 | 0.0575 | 0.052656 | 0.057834
| 1 | 5 | 0.06 | 0.0586 | 0.056 | 0.054 | 0.057 | 0.058 | 0.057 | 0.058 | 0.058 | 0.058 | 0.07837 | 0.06 | 0.057362 | 0.056 | 0.0556 | 0.050839 | 106
| 1 | 6 | 0.06 | 0.0578 | 0.058 | 0.054 | 0.057 | 0.058 | 0.057 | 0.059 | 0.06 | 0.057 | 0.0782 | 0.06 | 0.058266 | 0.057 | 0.0599 | 0.051415
| Mean |  | 0.06 | 0.0581 | 0.057167 | 0.054333 | 0.056833 | 0.057833 | 0.057167 | 0.059 | 0.057667 | 0.058167 | 0.077753 | 0.06 | 0.05671 | 0.058833 | 0.057867 | 0.050812
| Median |  | 0.06 | 0.0582 | 0.0575 | 0.054 | 0.057 | 0.058 | 0.057 | 0.059 | 0.0575 | 0.058 | 0.07819 | 0.06 | 0.057043 | 0.059 | 0.05785 | 0.051062
| Std Dev. |  | 0 | 0.00066633 | 0.00098319 | 0.0008165 | 0.00040825 | 0.00040825 | 0.00075277 | 0.00089443 | 0.0013663 | 0.001169 | 0.0012603 | 0 | 0.0012089 | 0.0020412 | 0.0017512 | 0.0013058
| Rel.Std.Dev. |  | 0% | 1.147% | 1.72% | 1.503% | 0.7183% | 0.7059% | 1.317% | 1.516% | 2.369% | 2.01% | 1.621% | 0% | 2.132% | 3.47% | 3.026% | 2.57%
| PDM3 |  | 3.745% | 0.4599% | -1.154% | -6.053% | -1.73% | -0.001178% | -1.154% | 2.016% | -0.2894% | 0.5752% | 34.44% | 3.745% | -1.944% | 1.728% | 0.05646% | -12.14%
| Z-Score (Absolute) |  | 1.45 | 0.17 | 0.46 | 2.37 | 0.69 | 0.01 | 0.46 | 0.78 | 0.12 | 0.21 | 13.42 | 1.45 | 0.77 | 0.66 | 0.01 | 4.75
| Analytical results for Pb in OREAS 143 (Certified Value 27.1 ppm) | [17 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES |  |  | 1
2
| 1 | 1 | 29.99 | 26 | 22 | 26.5 | 27.6 | 26.7 | 27.4 | 26.9 | 26.732 | 35 | 27.6505 | 35.6 | 23 | 22.22 |  |  | 1
| 1 | 2 | 29.89 | 27 | 20 | 27.7 | 26.8 | 27.7 | 26.7 | 29.4 | 26.613 | 33 | 27.6545 | 34.5 | 23 | 27.3 |  |  | 6
| 1 | 3 | 30.98 | 27 | 19 | 29 | 30.3 | 26.8 | 26.3 | 26 | 27.321 | 39 | 28.327 | 33 | 33 | 28.54 |  |  | 16
| 1 | 4 | 30.22 | 27 | 23 | 28.9 | 25 | 26.2 | 27.6 | 26.8 | 28.151 | 31 | 29.0039 | 40.6 | 46 | 31.82 |  |  | 27.0692
| 1 | 5 | 30.53 | 26 | 21 | 27.2 | 26.4 | 26.3 | 32.3 | 27.2 | 27.25 | 31 | 27.5629 | 34.4 | 25 | 26.83 |  |  | 107
| 1 | 6 | 31.06 | 26 | 19 | 27.4 | 28.6 | 26.4 | 27.4 | 27.9 | 28.908 | 29 | 28.1665 | 33.7 | 28 | 28.42
| Mean |  | 30.445 | 26.5 | 20.6667 | 27.7833 | 27.45 | 26.6833 | 27.95 | 27.3667 | 27.4958 | 33 | 28.0609 | 35.3 | 29.6667 | 27.5217
| Median |  | 30.375 | 26.5 | 20.5 | 27.55 | 27.2 | 26.55 | 27.4 | 27.05 | 27.2855 | 32 | 27.9105 | 34.45 | 26.5 | 27.86
| Std Dev. |  | 0.49746 | 0.54772 | 1.633 | 0.98675 | 1.8436 | 0.54924 | 2.1879 | 1.1708 | 0.8806 | 3.5777 | 0.55729 | 2.7379 | 8.8468 | 3.1291
| Rel.Std.Dev. |  | 1.634% | 2.067% | 7.902% | 3.552% | 6.716% | 2.058% | 7.828% | 4.278% | 3.203% | 10.84% | 1.986% | 7.756% | 29.82% | 11.37%
| PDM3 |  | 12.47% | -2.103% | -23.65% | 2.638% | 1.407% | -1.425% | 3.254% | 1.099% | 1.576% | 21.91% | 3.664% | 30.41% | 9.596% | 1.672%
| Z-Score (Absolute) |  | 2.73 | 1.13 | 6.84 | 0.13 | 0.2 | 0.95 | 0.29 | 0.28 | 0.15 | 5.24 | 0.4 | 7.49 | 1.97 | 0.13
| Analytical results for Pr in OREAS 143 (Certified Value 8.04 ppm) | [17 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 7.33 | 6.35 | 7.4 | 8.2 | 8.7852 | 7.7758 | 8.2859 | [9 empty] | 1
| 1 | 2 | 8.02 | 6.44 | 7.63 | 8.3 | 9.0469 | 7.9421 | 8.3803 | [9 empty] | #N/A
| 1 | 3 | 8.32 | 6.12 | 7.71 | 8.3 | 9.0199 | 8.1025 | 7.9694 | [9 empty] | 16
| 1 | 4 | 7.6 | 6.24 | 7.02 | 8.4 | 8.8662 | 8.0198 | 8.0056 | [9 empty] | 8.0388
| 1 | 5 | 7.24 | 6.64 | 7.78 | 8 | 8.0652 | 7.7459 | 7.9511 | [9 empty] | 108
| 1 | 6 | 7.63 | 6.76 | 7.53 | 8.1 | 8.6393 | 7.9917 | 8.2946
| Mean |  | 7.69 | 6.425 | 7.5117 | 8.2167 | 8.7371 | 7.9296 | 8.1478
| Median |  | 7.615 | 6.395 | 7.58 | 8.25 | 8.8257 | 7.9669 | 8.1457
| Std Dev. |  | 0.41221 | 0.24147 | 0.27564 | 0.1472 | 0.36221 | 0.14103 | 0.19257
| Rel.Std.Dev. |  | 5.36% | 3.758% | 3.669% | 1.791% | 4.146% | 1.779% | 2.363%
| PDM3 |  | -4.339% | -20.08% | -6.558% | 2.212% | 8.687% | -1.358% | 1.356%
| Z-Score (Absolute) |  | 0.56 | 3.53 | 0.98 | 0.67 | 1.9 | 0 | 0.51
| Analytical results for Rb in OREAS 143 (Certified Value 146 ppm) | [17 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 151.2 | 144.17 | 143 | 147 | 141 | 138 | 161 | 154.5 | 147 | 148.9469 | 172 | 145.4036 | 142 | 132.9508 |  |  | 1
| 1 | 2 | 147.9 | 144.33 | 139 | 147 | 143.5 | 141 | 161.5 | 160.5 | 149 | 148.534 | 169 | 145.002 | 139 | 130.6647 |  |  | 7
| 1 | 3 | 154.9 | 140.97 | 144 | 148 | 149 | 135 | 160 | 159 | 144 | 148.0797 | 187 | 145.7981 | 136 | 130.7363 |  |  | 16
| 1 | 4 | 154.4 | 135.72 | 140 | 152 | 152.5 | 141.5 | 152.5 | 160 | 147 | 148.1784 | 171 | 148.9539 | 132 | 130.0879 |  |  | 146.0925
| 1 | 5 | 154.5 | 140.26 | 143 | 151 | 137.5 | 140.5 | 157.5 | 158.5 | 146 | 148.6903 | 168 | 146.0203 | 136 | 130.4836 |  |  | 109
| 1 | 6 | 153.4 | 144.24 | 143 | 151 | 143.5 | 138.5 | 160 | 156 | 153.5 | 148.5876 | 166 | 151.0665 | 142 | 137.0402
| Mean |  | 152.7167 | 141.615 | 142 | 149.3333 | 144.5 | 139.0833 | 158.75 | 158.0833 | 147.75 | 148.5028 | 172.1667 | 147.0407 | 137.8333 | 131.9939
| Median |  | 153.9 | 142.57 | 143 | 149.5 | 143.5 | 139.5 | 160 | 158.75 | 147 | 148.5608 | 170 | 145.9092 | 137.5 | 130.7005
| Std Dev. |  | 2.711 | 3.3998 | 2 | 2.2509 | 5.4314 | 2.4376 | 3.3578 | 2.3541 | 3.2519 | 0.32397 | 7.5741 | 2.4203 | 3.92 | 2.6699
| Rel.Std.Dev. |  | 1.775% | 2.401% | 1.408% | 1.507% | 3.759% | 1.753% | 2.115% | 1.489% | 2.201% | 0.2182% | 4.399% | 1.646% | 2.844% | 2.023%
| PDM3 |  | 4.534% | -3.065% | -2.801% | 2.218% | -1.09% | -4.798% | 8.664% | 8.208% | 1.135% | 1.65% | 17.85% | 0.6491% | -5.653% | -9.65%
| Z-Score (Absolute) |  | 0.64 | 0.7 | 0.65 | 0.23 | 0.35 | 1 | 1.37 | 1.29 | 0.04 | 0.13 | 2.99 | 0.04 | 1.15 | 1.86
| Analytical results for Re in OREAS 143 (Indicative Value < 0.002 ppm) | [17 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 14 | 16 | 17 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 |  | < 0.002 | < 0.05 | < 0.001 | 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.01 | 0.002 | < 0.003 |  |  |  |  | 1
| 1 | 2 |  | < 0.002 | < 0.05 | < 0.001 | 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.01 | 0.003 | 0.034 |  |  |  |  | 18
| 1 | 3 | 0.03 | < 0.002 | < 0.05 | < 0.001 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.01 | 0.002 | < 0.003 |  |  |  |  | 16
| 1 | 4 |  | < 0.002 | < 0.05 | < 0.001 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.01 | 0.005 | < 0.003 |  |  |  |  | < 0.002
| 1 | 5 |  | < 0.002 | < 0.05 | < 0.001 | 0.002 | 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.01 | 0.003 | < 0.003 |  |  |  |  | 27
| 1 | 6 |  | < 0.002 | < 0.05 | < 0.001 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.01 | 0.004 | < 0.003
| Mean |  | 0.03 |  |  |  | 0.002 | 0.002 | 0.002 | 0.003 |  |  | 0.0031667 | 0.034
| Median |  | 0.03 |  |  |  | 0.002 | 0.002 | 0.002 | 0.003 |  |  | 0.003 | 0.034
| Std Dev. | [5 empty] | 0 | 0 |  |  |  |  | 0.001169
| Rel.Std.Dev. | [5 empty] | 0% | 0% |  |  |  |  | 36.92%
PDM3
| Z-Score (Absolute) |  | 20.84 | 0.49 | 17.16 | 0.86 | 0.25 | 0.25 | 0.37 | 0.25 | 0.49 | 2.45 | 1.1 | 3.86
| Analytical results for S in OREAS 143 (Certified Value 1.22 wt.%) | [17 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 1.2 | 1.16 | 1.22 | 1.17 | 1.23 | 1.24 | 1.19 | 1.26 | 1.25 | 1.296 | 1.1815 | 1.03 | 1.2045 | 1.24 | 1.2182 | 1.379 | 1
| 1 | 2 | 1.18 | 1.18 | 1.21 | 1.19 | 1.22 | 1.22 | 1.22 | 1.27 | 1.29 | 1.292 | 1.1728 | 1.01 | 1.2246 | 1.23 | 1.2236 | 1.366 | 24
| 1 | 3 | 1.22 | 1.17 | 1.25 | 1.19 | 1.23 | 1.22 | 1.2 | 1.26 | 1.21 | 1.294 | 1.1799 | 1.08 | 1.2238 | 1.23 | 1.2302 | 1.3034 | 16
| 1 | 4 | 1.2 | 1.17 | 1.24 | 1.24 | 1.24 | 1.21 | 1.22 | 1.31 | 1.22 | 1.305 | 1.1623 | 1.05 | 1.2668 | 1.24 | 1.2491 | 1.3843 | 1.2245
| 1 | 5 | 1.21 | 1.17 | 1.23 | 1.19 | 1.22 | 1.25 | 1.2 | 1.24 | 1.23 | 1.276 | 1.1733 | 1.01 | 1.2288 | 1.23 | 1.2226 | 1.3392 | 110
| 1 | 6 | 1.2 | 1.16 | 1.25 | 1.22 | 1.25 | 1.22 | 1.19 | 1.26 | 1.28 | 1.3 | 1.1682 | 1.01 | 1.2452 | 1.26 | 1.2209 | 1.4192
| Mean |  | 1.2017 | 1.1683 | 1.2333 | 1.2 | 1.2317 | 1.2267 | 1.2033 | 1.2667 | 1.2467 | 1.2938 | 1.173 | 1.0317 | 1.2323 | 1.2383 | 1.2274 | 1.3652
| Median |  | 1.2 | 1.17 | 1.235 | 1.19 | 1.23 | 1.22 | 1.2 | 1.26 | 1.24 | 1.295 | 1.173 | 1.02 | 1.2267 | 1.235 | 1.2231 | 1.3725
| Std Dev. |  | 0.013292 | 0.0075277 | 0.01633 | 0.025298 | 0.01169 | 0.015055 | 0.013663 | 0.023381 | 0.03266 | 0.0098877 | 0.0071582 | 0.028577 | 0.021306 | 0.01169 | 0.011341 | 0.039922
| Rel.Std.Dev. |  | 1.106% | 0.6443% | 1.324% | 2.108% | 0.9492% | 1.227% | 1.135% | 1.846% | 2.62% | 0.7642% | 0.6102% | 2.77% | 1.729% | 0.944% | 0.924% | 2.924%
| PDM3 |  | -1.866% | -4.588% | 0.7202% | -2.002% | 0.5841% | 0.1757% | -1.73% | 3.442% | 1.809% | 5.661% | -4.207% | -15.75% | 0.6337% | 1.128% | 0.2383% | 11.49%
| Z-Score (Absolute) |  | 0.7 | 1.53 | 0.09 | 0.74 | 0.05 | 0.07 | 0.65 | 0.93 | 0.43 | 1.6 | 1.41 | 4.93 | 0.07 | 0.22 | 0.05 | 3.38
| Analytical results for Sb in OREAS 143 (Certified Value 1.54 ppm) | [17 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES |  | 2
2
| 1 | 1 | 2.7 | 0.96 | 1.5 | 1.1 | 1.65 | 1.6 | 1.65 | 1.73 | 1.56 | < 5 | 1.73 | 1.5979 | 2.93 | 1.38 | 6.92 |  | 1
| 1 | 2 | 2.7 | 0.94 | 1.6 | 1.3 | 1.64 | 1.58 | 1.68 | 1.69 | 1.66 | < 5 | 1.41 | 1.7068 | 2.71 | 1.86 |  |  | 8
| 1 | 3 | 2.9 | 0.91 | 1.6 | 1.2 | 1.66 | 5.51 | 1.58 | 1.66 | 1.49 | < 5 | 1.7 | 1.548 | 3.12 | 1.38 | 7.07 |  | 16
| 1 | 4 | 3 | 0.89 | 1.5 | 1.2 | 1.76 | 1.6 | 1.47 | 1.76 | 1.54 | < 5 | 1.55 | 1.652 | 5.15 | 4.29 | 10.88 |  | 1.5409
| 1 | 5 | 3.7 | 0.92 | 1.6 | 1 | 1.68 | 1.6 | 1.58 | 2.14 | 1.59 | < 5 | 1.3 | 1.6287 | < 0.05 | 1.03 | 5.92 |  | 111
| 1 | 6 | 2.3 | 0.93 | 1.6 | 1.3 | 1.62 | 1.6 | 1.63 | 1.7 | 1.57 | 6.584 | 1.41 | 1.6785 | 5.04 | 1.19 | 4.6
| Mean |  | 2.8833 | 0.925 | 1.5667 | 1.1833 | 1.6683 | 2.2483 | 1.5983 | 1.78 | 1.5683 | 6.584 | 1.5167 | 1.6353 | 3.79 | 1.855 | 7.078
| Median |  | 2.8 | 0.925 | 1.6 | 1.2 | 1.655 | 1.6 | 1.605 | 1.715 | 1.565 | 6.584 | 1.48 | 1.6403 | 3.12 | 1.38 | 6.92
| Std Dev. |  | 0.46655 | 0.02429 | 0.05164 | 0.1169 | 0.04916 | 1.5979 | 0.07414 | 0.17967 | 0.056362 |  | 0.17317 | 0.05714 | 1.2007 | 1.225 | 2.3432
| Rel.Std.Dev. |  | 16.18% | 2.626% | 3.296% | 9.879% | 2.947% | 71.07% | 4.639% | 10.09% | 3.594% |  | 11.42% | 3.494% | 31.68% | 66.04% | 33.1%
| PDM3 |  | 87.12% | -39.97% | 1.672% | -23.21% | 8.27% | 45.91% | 3.727% | 15.52% | 1.78% | 327.3% | -1.573% | 6.128% | 146% | 20.38% | 359.3%
| Z-Score (Absolute) |  | 4.39 | 2.69 | 0.37 | 1.75 | 0 | 2.1 | 0.25 | 0.4 | 0.36 | 5.46 | 0.55 | 0.12 | 5.4 | 0.67 | 19.54
| Analytical results for Sc in OREAS 143 (Certified Value 10.1 ppm) | [17 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 10.39 | 10.03 | 10 | 10 | 9.2 | 10 | 10.3 | 10.4 | 9.7 | 10.407 | 9.6977 | 12.1 | 10.3722 | 8.4 | 10.0349 |  | 1
| 1 | 2 | 10.28 | 10.4 | 10 | 10 | 9.5 | 9.5 | 10.5 | 10.8 | 9.7 | 10.937 | 9.6941 | 11.5 | 10.549 | 8.2 | 9.8928 |  | 1
| 1 | 3 | 9.98 | 10.32 | 10 | 10 | 9.9 | 9.7 | 10.3 | 10.6 | 9.6 | 10.551 | 9.4486 | 12.6 | 10.517 | 8.2 | 9.7504 |  | 16
| 1 | 4 | 10.34 | 9.95 | 10 | 10 | 10 | 9.8 | 9.6 | 10.8 | 9.6 | 11.04 | 9.7415 | 11.8 | 10.3591 | 7.6 | 9.8314 |  | 10.108
| 1 | 5 | 9.89 | 10.2 | 10 | 10 | 9.2 | 10.2 | 10.1 | 10.5 | 9.8 | 10.719 | 9.7561 | 12 | 10.3911 | 7.8 | 9.8593 |  | 112
| 1 | 6 | 10.76 | 9.87 | 10 | 10 | 9.3 | 10 | 10.3 | 10.2 | 10.2 | 10.773 | 9.857 | 11.9 | 10.7906 | 7.8 | 10.4454
| Mean |  | 10.2733 | 10.1283 | 10 | 10 | 9.5167 | 9.8667 | 10.1833 | 10.55 | 9.7667 | 10.7378 | 9.6992 | 11.9833 | 10.4965 | 8 | 9.969
| Median |  | 10.31 | 10.115 | 10 | 10 | 9.4 | 9.9 | 10.3 | 10.55 | 9.7 | 10.746 | 9.7196 | 11.95 | 10.454 | 8 | 9.876
| Std Dev. |  | 0.31239 | 0.2116 | 0 | 0 | 0.35449 | 0.25033 | 0.31252 | 0.23452 | 0.22509 | 0.23534 | 0.13621 | 0.3656 | 0.16439 | 0.30984 | 0.25134
| Rel.Std.Dev. |  | 3.041% | 2.089% | 0% | 0% | 3.725% | 2.537% | 3.069% | 2.223% | 2.305% | 2.192% | 1.404% | 3.051% | 1.566% | 3.873% | 2.521%
| PDM3 |  | 1.636% | 0.2016% | -1.068% | -1.068% | -5.85% | -2.387% | 0.7457% | 4.373% | -3.376% | 6.232% | -4.044% | 18.55% | 3.844% | -20.85% | -1.374%
| Z-Score (Absolute) |  | 0.27 | 0 | NA | NA | 1.12 | 0.48 | 0.1 | 0.77 | 0.66 | 1.12 | 0.79 | 3.4 | 0.67 | 3.9 | 0.29
Results from laboratories 4 and 5 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 143 (Indicative Value 1.7 ppm) | [17 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 | 17 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2.16 | < 1 | 0.4 | 1 | 1 | 1 | < 1 | < 1 | 4 | < 1.5 | 2 | 2.61 |  |  |  |  | 1
| 1 | 2 | 2.18 | < 1 | 0.4 | 1 | 1 | < 1 | 1 | < 1 | 3 | < 1.5 | 2 | 1.63 |  |  |  |  | 22
| 1 | 3 | 2.04 | < 1 | 0.3 | 1 | 1 | < 1 | < 1 | < 1 | < 2 | < 1.5 | 2 | 2.77 |  |  |  |  | 16
| 1 | 4 | 2.16 | < 1 | 0.3 | 1 | 1 | < 1 | < 1 | < 1 | 3 | < 1.5 | 2 | 4.14 |  |  |  |  | 1.7011
| 1 | 5 | 2.05 | < 1 | 0.3 | 2 | 1 | < 1 | 1 | < 1 | < 2 | < 1.5 | 4 | 3.97 |  |  |  |  | 28
| 1 | 6 | 2.08 | < 1 | 0.4 | 1 | 1 | < 1 | < 1 | < 1 | < 2 | < 1.5 | 1 | 3.97
| Mean |  | 2.1117 |  | 0.35 | 1.1667 | 1 | 1 | 1 |  | 3.3333 |  | 2.1667 | 3.1817
| Median |  | 2.12 |  | 0.35 | 1 | 1 | 1 | 1 |  | 3 |  | 2 | 3.37
| Std Dev. |  | 0.062102 |  | 0.054772 | 0.40825 | 0 |  | 0 |  | 0.57735 |  | 0.98319 | 1.0065
| Rel.Std.Dev. |  | 2.941% |  | 15.65% | 34.99% | 0% |  | 0% |  | 17.32% |  | 45.38% | 31.63%
| PDM3 |  | 24.13% |  | -79.43% | -31.42% | -41.21% | -41.21% | -41.21% |  | 95.95% |  | 27.37% | 87.03%
| Z-Score (Absolute) |  | 2.22 | 0.67 | 0.94 | 0.52 | 0.22 | 0.52 | 0.37 | 0.67 | 2.32 | 0.22 | 2.32 | 4.15
| Analytical results for Sm in OREAS 143 (Certified Value 5.65 ppm) | [17 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 5.5 | 5.25 | 5.44 | 6.3 | 5.1581 | 6.0934 | 5.3217 | [9 empty] | 1
| 1 | 2 | 6 | 5.6 | 5.83 | 5.9 | 5.2542 | 5.8551 | 5.3055 | [9 empty] | #N/A
| 1 | 3 | 6.2 | 5.24 | 5.89 | 5.7 | 5.2839 | 6.0669 | 5.2263 | [9 empty] | 16
| 1 | 4 | 5.9 | 5.33 | 5.37 | 5.6 | 5.4426 | 6.1152 | 5.3429 | [9 empty] | 5.6545
| 1 | 5 | 5.6 | 5.31 | 5.83 | 6 | 5.5059 | 5.8495 | 5.5442 | [9 empty] | 113
| 1 | 6 | 5.9 | 5.44 | 5.77 | 6.3 | 5.1669 | 6.1197 | 5.635
| Mean |  | 5.85 | 5.3617 | 5.6883 | 5.9667 | 5.3019 | 6.0167 | 5.3959
| Median |  | 5.9 | 5.32 | 5.8 | 5.95 | 5.269 | 6.0802 | 5.3323
| Std Dev. |  | 0.25884 | 0.13703 | 0.22382 | 0.29439 | 0.14347 | 0.12869 | 0.15774
| Rel.Std.Dev. |  | 4.425% | 2.556% | 3.935% | 4.934% | 2.706% | 2.139% | 2.923%
| PDM3 |  | 3.458% | -5.178% | 0.5992% | 5.522% | -6.235% | 6.406% | -4.572%
| Z-Score (Absolute) |  | 0.37 | 0.75 | 0 | 0.64 | 0.89 | 0.76 | 0.67
| Analytical results for Sn in OREAS 143 (Indicative Value 0.013 wt.%) | [17 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 17 | 18 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0076 | 0.01377 | 0.0156 | 0.0038 | 0.00326 | 0.00364 | 0.01615 | > 0.05 | 0.047 | 0.036 | 0.0548 | 0.0093318 | 0.00418 | 0.024152 | 0.25692 |  | 1
| 1 | 2 | 0.0067 | 0.01375 | 0.0153 | 0.0039 | 0.00491 | 0.00369 | 0.00895 | > 0.05 | 0.00985 | 0.034 | 0.0543 | 0.0082772 | 0.00985 | 0.020902 | 0.27469 |  | 20
| 1 | 3 | 0.0114 | 0.013389 | 0.0156 | 0.0038 | 0.00462 | 0.00346 | 0.00773 | 0.049 | > 0.05 | 0.034 | 0.0578 | 0.0091026 | 0.00684 | 0.021182 | 0.25481 |  | 16
| 1 | 4 | 0.009 | 0.013199 | 0.0155 | 0.0041 | 0.00418 | 0.00402 | 0.00329 | > 0.05 | 0.048 | 0.034 | 0.0534 | 0.011505 | 0.0109 | 0.021822 | 0.28002 |  | 0.013489
| 1 | 5 | 0.0121 | 0.012732 | 0.0161 | 0.0045 | 0.0045 | 0.00356 | 0.01205 | > 0.05 | 0.00807 | 0.037 | 0.0532 | 0.0085388 | 0.00604 | 0.018412 | 0.26902 |  | 26
| 1 | 6 | 0.0139 | 0.012773 | 0.015 | 0.0035 | 0.00349 | 0.00349 | 0.0322 | > 0.05 | 0.05 | 0.036 | 0.0523 | 0.0091204 | 0.0123 | 0.019402 | 0.25586
| Mean |  | 0.010117 | 0.013269 | 0.015517 | 0.0039333 | 0.00416 | 0.0036433 | 0.013395 | 0.049 | 0.032584 | 0.035167 | 0.0543 | 0.0093126 | 0.0083517 | 0.020979 | 0.26522
| Median |  | 0.0102 | 0.013294 | 0.01555 | 0.00385 | 0.00434 | 0.0036 | 0.0105 | 0.049 | 0.047 | 0.035 | 0.05385 | 0.0091115 | 0.008345 | 0.021042 | 0.26297
| Std Dev. |  | 0.0027982 | 0.0004552 | 0.0003656 | 0.00033862 | 0.00065559 | 0.00020403 | 0.010168 |  | 0.021602 | 0.0013292 | 0.0019246 | 0.0011453 | 0.003142 | 0.0019951 | 0.010843
| Rel.Std.Dev. |  | 27.66% | 3.431% | 2.356% | 8.609% | 15.76% | 5.6% | 75.91% |  | 66.3% | 3.78% | 3.544% | 12.3% | 37.62% | 9.51% | 4.088%
| PDM3 |  | -25% | -1.636% | 15.03% | -70.84% | -69.16% | -72.99% | -0.7005% | 263.2% | 141.6% | 160.7% | 302.5% | -30.96% | -38.09% | 55.52% | 1866%
| Z-Score (Absolute) |  | 0.24 | 0.01 | 0.15 | 0.69 | 0.67 | 0.71 | 0 | 2.6 | 1.4 | 1.59 | 2.99 | 0.3 | 0.37 | 0.55 | 18.39
| Analytical results for Sr in OREAS 143 (Certified Value 128 ppm) | [17 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  | 0
0
| 1 | 1 | 129 | 126.86 | 123 | 134 | 123.5 | 129.5 | 130.5 | 132 | 127.5 | 131.382 | 124.1896 | 130 | 123.4198 | 81 | 120.33 |  | 1
| 1 | 2 | 130.5 | 127.97 | 125 | 133 | 125.5 | 128.5 | 132.5 | 135 | 125.5 | 132.456 | 123.87 | 135 | 125.0135 | 79.6 | 124.29 |  | #N/A
| 1 | 3 | 131.4 | 124.14 | 124 | 136 | 128 | 130 | 132.5 | 133 | 125.5 | 131.606 | 120.7418 | 139 | 126.484 | 80.3 | 115.41 |  | 16
| 1 | 4 | 131.5 | 121.8 | 122 | 131 | 129 | 128 | 132 | 138 | 123 | 132.475 | 124.7585 | 132 | 129.5807 | 81 | 123.6 |  | 128.0104
| 1 | 5 | 128.7 | 124.32 | 121 | 131 | 124 | 131.5 | 127.5 | 133.5 | 126 | 130.889 | 122.1704 | 131 | 126.4653 | 79.6 | 118.39 |  | 114
| 1 | 6 | 134.1 | 123.8 | 123 | 132 | 126.5 | 131 | 128 | 132 | 133.5 | 131.671 | 126.0799 | 131 | 128.8883 | 81.7 | 116.32
| Mean |  | 130.8667 | 124.815 | 123 | 132.8333 | 126.0833 | 129.75 | 130.5 | 133.9167 | 126.8333 | 131.7465 | 123.635 | 133 | 126.6419 | 80.5333 | 119.7233
| Median |  | 130.95 | 124.23 | 123 | 132.5 | 126 | 129.75 | 131.25 | 133.25 | 125.75 | 131.6385 | 124.0298 | 131.5 | 126.4747 | 80.65 | 119.36
| Std Dev. |  | 1.9725 | 2.2342 | 1.4142 | 1.9408 | 2.1775 | 1.3693 | 2.2583 | 2.2895 | 3.573 | 0.62092 | 1.9036 | 3.4059 | 2.314 | 0.84774 | 3.6937
| Rel.Std.Dev. |  | 1.507% | 1.79% | 1.15% | 1.461% | 1.727% | 1.055% | 1.731% | 1.71% | 2.817% | 0.4713% | 1.54% | 2.561% | 1.827% | 1.053% | 3.085%
| PDM3 |  | 2.231% | -2.496% | -3.914% | 3.768% | -1.505% | 1.359% | 1.945% | 4.614% | -0.9195% | 2.919% | -3.418% | 3.898% | -1.069% | -37.09% | -6.474%
| Z-Score (Absolute) |  | 0.71 | 0.36 | 0.67 | 1.06 | 0.13 | 0.51 | 0.64 | 1.25 | 0 | 0.86 | 0.56 | 1.08 | 0.03 | 8.14 | 1.25
| Analytical results for Ta in OREAS 143 (Certified Value 0.96 ppm) | [17 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.9 | 1.1 | 1.13 | 0.78 | 0.7 | 0.85 | 0.95 | 0.91 | 0.96 | 0.9 | 0.97 | 0.93635 | 0.56 | 5.07 |  |  | 1
| 1 | 2 | 0.8 | 1.15 | 1.12 | 0.76 | 1 | 0.91 | 0.96 | 0.93 | 0.98 | 0.88 | 0.84 | 1.0258 | 0.55 | 4.14 |  |  | 3
| 1 | 3 | 1 | 1.17 | 1.12 | 0.83 | 0.8 | 0.91 | 0.88 | 0.94 | 0.98 | 0.87 | 0.99 | 1.0097 | 0.4 | 2.09 |  |  | 16
| 1 | 4 | 0.9 | 1.13 | 1.16 | 0.77 | 0.8 | 0.92 | 0.89 | 0.93 | 1.01 | 0.93 | 0.85 | 0.98421 | 0.21 | 4.1214 |  |  | 0.95618
| 1 | 5 | 1 | 1.14 | 1.14 | 0.76 | 0.8 | 0.88 | 0.89 | 0.88 | 0.97 | 0.92 | 0.85 | 0.9573 | 0.45 | 2.28 |  |  | 115
| 1 | 6 | 0.9 | 1.11 | 1.19 | 0.78 | 0.9 | 0.89 | 0.95 | 0.9 | 1 | 0.95 | 0.84 | 1.0577 | 0.6 | 3.5441
| Mean |  | 0.91667 | 1.1333 | 1.1433 | 0.78 | 0.83333 | 0.89333 | 0.92 | 0.915 | 0.98333 | 0.90833 | 0.89 | 0.99517 | 0.46167 | 3.5409
| Median |  | 0.9 | 1.135 | 1.135 | 0.775 | 0.8 | 0.9 | 0.92 | 0.92 | 0.98 | 0.91 | 0.85 | 0.99695 | 0.5 | 3.8327
| Std Dev. |  | 0.075277 | 0.02582 | 0.027325 | 0.026077 | 0.10328 | 0.02582 | 0.036878 | 0.022583 | 0.018619 | 0.030605 | 0.070143 | 0.044877 | 0.14414 | 1.1601
| Rel.Std.Dev. |  | 8.212% | 2.278% | 2.39% | 3.343% | 12.39% | 2.89% | 4.008% | 2.468% | 1.893% | 3.369% | 7.881% | 4.509% | 31.22% | 32.76%
| PDM3 |  | -4.133% | 18.53% | 19.57% | -18.43% | -12.85% | -6.573% | -3.784% | -4.307% | 2.839% | -5.004% | -6.922% | 4.077% | -51.72% | 270.3%
| Z-Score (Absolute) |  | NA | 2.03 | 2.12 | 1.29 | NA | 0.23 | 0.02 | 0.02 | 0.62 | 0.09 | 0.26 | 0.73 | 4.28 | 24.66
Results from laboratories 1 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 143 (Certified Value 0.6 ppm) | [17 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 13 | 14 | 18 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [8 empty] | 2
3
| 1 | 1 | 0.58 | 0.66 | 0.55 | 0.7 | 0.61922 | 0.6 | 0.60848 | 0.5711 | [8 empty] | 1
| 1 | 2 | 0.6 | 0.65 | 0.57 | 0.7 | 0.6109 | 0.6 | 0.60954 | 0.5645 | [8 empty] | #N/A
| 1 | 3 | 0.66 | 0.65 | 0.54 | 0.7 | 0.59313 | 0.64 | 0.61434 | 0.5687 | [8 empty] | 16
| 1 | 4 | 0.63 | 0.66 | 0.53 | 0.7 | 0.60808 | 0.6 | 0.6096 | 0.5681 | [8 empty] | 0.60025
| 1 | 5 | 0.61 | 0.65 | 0.56 | 0.7 | 0.53571 | 0.62 | 0.60313 | 0.5859 | [8 empty] | 116
| 1 | 6 | 0.59 | 0.66 | 0.54 | 0.7 | 0.57438 | 0.6 | 0.61765 | 0.5981
| Mean |  | 0.61167 | 0.655 | 0.54833 | 0.7 | 0.59024 | 0.61 | 0.61046 | 0.57607
| Median |  | 0.605 | 0.655 | 0.545 | 0.7 | 0.6006 | 0.6 | 0.60957 | 0.5699
| Std Dev. |  | 0.029269 | 0.0054772 | 0.01472 | 1.2162e-16 | 0.031042 | 0.016733 | 0.0050186 | 0.013102
| Rel.Std.Dev. |  | 4.785% | 0.8362% | 2.684% | 1.737e-14% | 5.259% | 2.743% | 0.8221% | 2.274%
| PDM3 |  | 1.902% | 9.121% | -8.649% | 16.62% | -1.669% | 1.624% | 1.7% | -4.029%
| Z-Score (Absolute) |  | 0.06 | 1.54 | 2.1 | NA | 0.67 | 0 | 0.02 | 1.16
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 143 (Indicative Value 0.032 ppm) | [17 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 | 17 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.43 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.2 | 0.27 | < 0.03 |  |  |  |  | 1
| 1 | 2 | 0.43 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | 0.05 | < 0.05 | < 0.2 | 0.56 | 0.5 |  |  |  |  | 4
| 1 | 3 | 0.44 | < 0.1 | < 0.1 | < 0.05 | 0.05 | 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.2 | < 0.01 | < 0.03 |  |  |  |  | 16
| 1 | 4 | 0.43 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | 0.05 | < 0.05 | < 0.2 | 0.41 | < 0.03 |  |  |  |  | 0.031944
| 1 | 5 | 0.44 | < 0.1 | < 0.1 | < 0.05 | 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.2 | 0.51 | 0.23 |  |  |  |  | 27
| 1 | 6 | 0.46 | < 0.1 | < 0.1 | < 0.05 | 0.05 | 0.07 | < 0.05 | 0.06 | < 0.05 | < 0.2 | 0.28 | < 0.03
| Mean |  | 0.43833 |  |  |  | 0.05 | 0.06 | 0.053333 | 0.055 | 0.05 |  | 0.406 | 0.365
| Median |  | 0.435 |  |  |  | 0.05 | 0.06 | 0.05 | 0.055 | 0.05 |  | 0.41 | 0.365
| Std Dev. |  | 0.01169 |  |  |  | 8.4984e-18 | 0.014142 | 0.0057735 | 0.0057735 |  |  | 0.13126 | 0.19092
| Rel.Std.Dev. |  | 2.667% |  |  |  | 1.7e-14% | 23.57% | 10.83% | 10.5% |  |  | 32.33% | 52.31%
| PDM3 |  | 1272% |  |  |  | 56.52% | 87.83% | 66.96% | 72.17% | 56.52% |  | 1171% | 1043%
| Z-Score (Absolute) |  | 18.07 | 0.12 | 0.12 | 1.04 | 0.46 | 0.5 | 0.39 | 0.12 | 0.85 | 2.43 | 13.49 | 3.89
| Analytical results for Th in OREAS 143 (Certified Value 12.9 ppm) | [17 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 12.1 | 11.63 | 12.14 | 12.5 | 12.4 | 12 | 12.45 | 13.55 | 12.85 | 12.9 | 12.9 | 13.7671 | 15.4 | 13.5407 |  |  | 1
| 1 | 2 | 12.9 | 11.89 | 12.24 | 12.7 | 12.6 | 12.55 | 13.1 | 13.85 | 13.2 | 13.15 | 13.4 | 13.7924 | 14.8 | 13.5853 |  |  | #N/A
| 1 | 3 | 13 | 11.34 | 12.44 | 12.4 | 12.5 | 13 | 12.05 | 13.7 | 13.3 | 12.85 | 13.8 | 14.4953 | 14.5 | 12.7148 |  |  | 16
| 1 | 4 | 12.6 | 11.42 | 12.19 | 12.6 | 13 | 12.9 | 11.85 | 13.05 | 13.6 | 12.95 | 13.3 | 14.4239 | 13.9 | 13.7215 |  |  | 12.9408
| 1 | 5 | 12.8 | 11.61 | 12.48 | 12.4 | 12.1 | 11.85 | 12.2 | 13.15 | 13.1 | 12.95 | 13.4 | 13.7685 | 14.1 | 13.0741 |  |  | 117
| 1 | 6 | 12.5 | 11.51 | 12.28 | 12.5 | 12.2 | 12.15 | 13.25 | 13.3 | 13 | 13.65 | 13.3 | 13.7035 | 14.3 | 14.0009
| Mean |  | 12.65 | 11.5667 | 12.295 | 12.5167 | 12.4667 | 12.4083 | 12.4833 | 13.4333 | 13.175 | 13.075 | 13.35 | 13.9918 | 14.5 | 13.4395
| Median |  | 12.7 | 11.56 | 12.26 | 12.5 | 12.45 | 12.35 | 12.325 | 13.425 | 13.15 | 12.95 | 13.35 | 13.7804 | 14.4 | 13.563
| Std Dev. |  | 0.32711 | 0.19315 | 0.13678 | 0.1169 | 0.32042 | 0.48106 | 0.57242 | 0.31728 | 0.26029 | 0.29958 | 0.2881 | 0.36426 | 0.54037 | 0.46581
| Rel.Std.Dev. |  | 2.586% | 1.67% | 1.113% | 0.934% | 2.57% | 3.877% | 4.585% | 2.362% | 1.976% | 2.291% | 2.158% | 2.603% | 3.727% | 3.466%
| PDM3 |  | -2.247% | -10.62% | -4.99% | -3.278% | -3.664% | -4.115% | -3.535% | 3.806% | 1.81% | 1.037% | 3.162% | 8.121% | 12.05% | 3.854%
| Z-Score (Absolute) |  | 0.3 | 1.86 | 0.81 | 0.5 | 0.57 | 0.65 | 0.54 | 0.82 | 0.45 | 0.3 | 0.7 | 1.62 | 2.35 | 0.83
| Analytical results for Ti in OREAS 143 (Certified Value 0.325 wt.%) | [17 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.2908 | 0.35 | 0.302 | 0.335 | 0.332 | 0.335 | 0.332 | 0.349 | 0.341 | 0.313 | 0.30329 | 0.3 | 0.32807 | 0.338 | 0.30198 |  | 1
| 1 | 2 | 0.2972 | 0.35 | 0.296 | 0.339 | 0.327 | 0.333 | 0.339 | 0.351 | 0.341 | 0.322 | 0.30749 | 0.31 | 0.33157 | 0.322 | 0.30598 |  | #N/A
| 1 | 3 | 0.3064 | 0.35 | 0.321 | 0.334 | 0.327 | 0.334 | 0.335 | 0.35 | 0.334 | 0.316 | 0.30811 | 0.33 | 0.33659 | 0.32 | 0.28901 |  | 16
| 1 | 4 | 0.2929 | 0.34 | 0.309 | 0.336 | 0.328 | 0.331 | 0.333 | 0.354 | 0.338 | 0.32 | 0.30596 | 0.32 | 0.34015 | 0.317 | 0.31291 |  | 0.32478
| 1 | 5 | 0.2908 | 0.35 | 0.304 | 0.332 | 0.329 | 0.339 | 0.324 | 0.346 | 0.339 | 0.318 | 0.30895 | 0.31 | 0.33324 | 0.319 | 0.29948 |  | 118
| 1 | 6 | 0.2964 | 0.35 | 0.308 | 0.336 | 0.333 | 0.334 | 0.33 | 0.347 | 0.356 | 0.321 | 0.30561 | 0.31 | 0.34048 | 0.333 | 0.295
| Mean |  | 0.29575 | 0.34833 | 0.30667 | 0.33533 | 0.32933 | 0.33433 | 0.33217 | 0.3495 | 0.3415 | 0.31833 | 0.30657 | 0.31333 | 0.33502 | 0.32483 | 0.30072
| Median |  | 0.29465 | 0.35 | 0.306 | 0.3355 | 0.3285 | 0.334 | 0.3325 | 0.3495 | 0.34 | 0.319 | 0.30672 | 0.31 | 0.33492 | 0.321 | 0.30073
| Std Dev. |  | 0.0058835 | 0.0040825 | 0.0084301 | 0.0023381 | 0.002582 | 0.0026583 | 0.0050365 | 0.002881 | 0.0075565 | 0.0033862 | 0.0020453 | 0.010328 | 0.0049374 | 0.0085654 | 0.0083588
| Rel.Std.Dev. |  | 1.989% | 1.172% | 2.749% | 0.6972% | 0.784% | 0.7951% | 1.516% | 0.8243% | 2.213% | 1.064% | 0.6672% | 3.296% | 1.474% | 2.637% | 2.78%
| PDM3 |  | -8.939% | 7.252% | -5.578% | 3.249% | 1.401% | 2.941% | 2.274% | 7.611% | 5.148% | -1.985% | -5.608% | -3.525% | 3.151% | 0.0159% | -7.407%
| Z-Score (Absolute) |  | 1.86 | 1.05 | 1.26 | 0.33 | 0 | 0.28 | 0.16 | 1.12 | 0.67 | 0.61 | 1.26 | 0.89 | 0.31 | 0.25 | 1.59
| Analytical results for Tl in OREAS 143 (Certified Value 0.79 ppm) | [17 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 16 | 17 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.8 | 0.83 | 0.79 | 0.79 | 0.78 | 0.72 | 0.81 | 0.76 | 0.79 | 0.79 | 0.76 | 0.82527 | 0.9 | 0.641 | 0.785 |  | 1
| 1 | 2 | 0.9 | 0.84 | 0.8 | 0.77 | 0.8 | 0.83 | 0.79 | 0.82 | 0.81 | 0.79 | 0.77 | 0.81992 | 0.86 | 0.592 | 0.7656 |  | 5
| 1 | 3 | 0.8 | 0.85 | 0.81 | 0.81 | 0.8 | 0.8 | 0.76 | 0.81 | 0.82 | 0.78 | 0.83 | 0.83408 | 0.84 | 0.53 | 0.7551 |  | 16
| 1 | 4 | 0.8 | 0.83 | 0.78 | 0.78 | 0.79 | 0.82 | 0.76 | 0.73 | 0.8 | 0.82 | 0.81 | 0.87877 | 0.83 | 0.681 | 0.755 |  | 0.78852
| 1 | 5 | 0.9 | 0.86 | 0.78 | 0.77 | 0.8 | 0.73 | 0.73 | 0.77 | 0.8 | 0.8 | 0.76 | 0.80096 | 0.8 | 0.674 | 0.7568 |  | 119
| 1 | 6 | 0.8 | 0.86 | 0.81 | 0.78 | 0.81 | 0.77 | 0.83 | 0.78 | 0.81 | 0.8 | 0.79 | 0.80225 | 0.87 | 0.649 | 0.7534
| Mean |  | 0.83333 | 0.845 | 0.795 | 0.78333 | 0.79667 | 0.77833 | 0.78 | 0.77833 | 0.805 | 0.79667 | 0.78667 | 0.82687 | 0.85 | 0.62783 | 0.76182
| Median |  | 0.8 | 0.845 | 0.795 | 0.78 | 0.8 | 0.785 | 0.775 | 0.775 | 0.805 | 0.795 | 0.78 | 0.82259 | 0.85 | 0.645 | 0.75595
| Std Dev. |  | 0.05164 | 0.013784 | 0.013784 | 0.015055 | 0.010328 | 0.046224 | 0.036878 | 0.033116 | 0.010488 | 0.013663 | 0.028752 | 0.028545 | 0.034641 | 0.057332 | 0.01216
| Rel.Std.Dev. |  | 6.197% | 1.631% | 1.734% | 1.922% | 1.296% | 5.939% | 4.728% | 4.255% | 1.303% | 1.715% | 3.655% | 3.452% | 4.075% | 9.132% | 1.596%
| PDM3 |  | 5.684% | 7.163% | 0.8223% | -0.6573% | 1.034% | -1.291% | -1.08% | -1.291% | 2.091% | 1.034% | -0.2345% | 4.865% | 7.797% | -20.38% | -3.386%
| Z-Score (Absolute) |  | NA | 2.92 | 0.22 | 0.4 | 0.31 | 0.67 | 0.58 | 0.67 | 0.76 | 0.31 | 0.22 | 1.94 | 3.19 | 8.79 | 1.57
Results from laboratory 1 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 143 (Certified Value 0.18 ppm) | [17 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 14 | 18 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 0.2 | < 1 | 0.17 | 0.2 | 0.17108 | 0.23955 | 0.1612 | [9 empty] | 1
| 1 | 2 | 0.2 | < 1 | 0.17 | 0.2 | 0.15289 | 0.23516 | 0.1587 | [9 empty] | #N/A
| 1 | 3 | 0.2 | < 1 | 0.17 | 0.2 | 0.17365 | 0.24173 | 0.1633 | [9 empty] | 16
| 1 | 4 | 0.2 | < 1 | 0.17 | 0.2 | 0.17456 | 0.24265 | 0.1587 | [9 empty] | 0.17964
| 1 | 5 | 0.1 | < 1 | 0.17 | 0.2 | 0.15453 | 0.22132 | 0.1674 | [9 empty] | 120
| 1 | 6 | 0.2 | < 1 | 0.17 | 0.2 | 0.1534 | 0.23468 | 0.1797
| Mean |  | 0.18333 |  | 0.17 | 0.2 | 0.16335 | 0.23585 | 0.16483
| Median |  | 0.2 |  | 0.17 | 0.2 | 0.16281 | 0.23736 | 0.16225
| Std Dev. |  | 0.040825 |  | 0 | 3.0405e-17 | 0.010747 | 0.0078393 | 0.0079799
| Rel.Std.Dev. |  | 22.27% |  | 0% | 1.52e-14% | 6.579% | 3.324% | 4.841%
| PDM3 |  | 2.058% |  | -5.365% | 11.34% | -9.065% | 31.29% | -8.241%
| Z-Score (Absolute) |  | 0 | 11.54 | 0.49 | 0.61 | 0.73 | 1.91 | 0.67
| Analytical results for U in OREAS 143 (Certified Value 2.8 ppm) | [17 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 3.1 | 2.55 | 2.88 | 2.7 | 3 | 2.6 | 2.7 | 2.7 | 2.6 | 2.7 | 2.5059 | 2.96 | 2.9357 | 2.94 | 2.44 | 2.8452 | 1
| 1 | 2 | 2.6 | 2.85 | 2.77 | 2.83 | 3.2 | 2.5 | 2.7 | 2.8 | 3 | 2.8 | 2.5913 | 3.27 | 3.3186 | 3.04 | 2.32 | 2.7291 | #N/A
| 1 | 3 | 2.8 | 2.7 | 2.65 | 2.78 | 2.8 | 2.8 | 2.7 | 2.7 | 3.3 | 2.7 | 2.5386 | 3.14 | 3.1275 | 3.42 | 2.11 | 3.0953 | 16
| 1 | 4 | 2.7 | 2.62 | 2.84 | 2.72 | 3.1 | 2.6 | 2.8 | 2.7 | 3.1 | 2.7 | 2.8963 | 3.18 | 3.2955 | 2.67 | 2.71 | 3.0678 | 2.8044
| 1 | 5 | 2.6 | 2.77 | 2.84 | 2.73 | 3 | 2.7 | 3.1 | 2.8 | 2.6 | 2.9 | 2.4208 | 3.04 | 2.8674 | 3.08 | 2.17 | 2.7695 | 121
| 1 | 6 | 2.7 | 2.6 | 2.8 | 2.76 | 3 | 2.6 | 2.8 | 2.6 | 2.6 | 3 | 2.1616 | 3.13 | 3.0878 | 2.83 | 2.57 | 3.0569
| Mean |  | 2.75 | 2.6817 | 2.7967 | 2.7533 | 3.0167 | 2.6333 | 2.8 | 2.7167 | 2.8667 | 2.8 | 2.5191 | 3.12 | 3.1054 | 2.9967 | 2.3867 | 2.9273
| Median |  | 2.7 | 2.66 | 2.82 | 2.745 | 3 | 2.6 | 2.75 | 2.7 | 2.8 | 2.75 | 2.5223 | 3.135 | 3.1077 | 2.99 | 2.38 | 2.9511
| Std Dev. |  | 0.18708 | 0.11339 | 0.081158 | 0.047188 | 0.13292 | 0.10328 | 0.15492 | 0.075277 | 0.30768 | 0.12649 | 0.23895 | 0.10826 | 0.18322 | 0.25524 | 0.23192 | 0.16473
| Rel.Std.Dev. |  | 6.803% | 4.228% | 2.902% | 1.714% | 4.406% | 3.922% | 5.533% | 2.771% | 10.73% | 4.518% | 9.486% | 3.47% | 5.9% | 8.517% | 9.717% | 5.627%
| PDM3 |  | -1.939% | -4.376% | -0.2752% | -1.82% | 7.57% | -6.099% | -0.1564% | -3.128% | 2.221% | -0.1564% | -10.17% | 11.25% | 10.73% | 6.856% | -14.9% | 4.383%
| Z-Score (Absolute) |  | 0.27 | 0.64 | 0.01 | 0.25 | 1.2 | 0.91 | 0.01 | 0.45 | 0.38 | 0.01 | 1.53 | 1.77 | 1.69 | 1.09 | 2.26 | 0.71
| Analytical results for V in OREAS 143 (Certified Value 72 ppm) | [17 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 0
1
| 1 | 1 | 71 | 71 | 73 | 75 | 71 | 72 | 70 | 73 | 72 | 71.545 | 72.0986 | 72 | 69.968 | 72 | 108 | 70.081 | 1
| 1 | 2 | 72 | 71 | 73 | 77 | 71 | 72 | 72 | 72 | 71 | 72.911 | 70.3632 | 72 | 71.4901 | 71 | 107 | 70.631 | #N/A
| 1 | 3 | 73 | 70 | 73 | 74 | 71 | 71 | 70 | 73 | 70 | 72.189 | 70.5788 | 78 | 72.6816 | 74 | 108 | 68.057 | 16
| 1 | 4 | 70 | 69 | 73 | 77 | 71 | 73 | 72 | 73 | 71 | 73.413 | 72.25 | 74 | 74.0348 | 70 | 110 | 71.533 | 71.734
| 1 | 5 | 71 | 70 | 73 | 74 | 71 | 72 | 70 | 72 | 72 | 71.65 | 71.7548 | 71 | 71.4827 | 68 | 108 | 69.146 | 122
| 1 | 6 | 71 | 71 | 73 | 74 | 71 | 72 | 70 | 72 | 74 | 71.82 | 71.549 | 75 | 72.8363 | 67 | 108 | 67.199
| Mean |  | 71.3333 | 70.3333 | 73 | 75.1667 | 71 | 72 | 70.6667 | 72.5 | 71.6667 | 72.2547 | 71.4324 | 73.6667 | 72.0823 | 70.3333 | 108.1667 | 69.4412
| Median |  | 71 | 70.5 | 73 | 74.5 | 71 | 72 | 70 | 72.5 | 71.5 | 72.0045 | 71.6519 | 73 | 72.0859 | 70.5 | 108 | 69.6135
| Std Dev. |  | 1.0328 | 0.8165 | 0 | 1.472 | 0 | 0.63246 | 1.0328 | 0.54772 | 1.3663 | 0.75299 | 0.78761 | 2.582 | 1.4086 | 2.582 | 0.98319 | 1.6267
| Rel.Std.Dev. |  | 1.448% | 1.161% | 0% | 1.958% | 0% | 0.8784% | 1.462% | 0.7555% | 1.906% | 1.042% | 1.103% | 3.505% | 1.954% | 3.671% | 0.909% | 2.343%
| PDM3 |  | -0.5586% | -1.953% | 1.765% | 4.785% | -1.023% | 0.3708% | -1.488% | 1.068% | -0.09391% | 0.7258% | -0.4205% | 2.694% | 0.4854% | -1.953% | 50.79% | -3.196%
| Z-Score (Absolute) |  | 0.34 | 1.01 | 0.79 | 2.25 | 0.56 | 0.11 | 0.79 | 0.45 | 0.11 | 0.28 | 0.27 | 1.24 | 0.17 | 1.01 | 24.5 | 1.61
| Analytical results for W in OREAS 143 (Certified Value 21.3 ppm) | [17 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 20 | 7.98 | 21.1 | 19.1 | 20.8 | 24.3 | 21.2 | 21.7 | 24.1 | 30.569 | 15.75 | 20 | 21.1189 | 21.7 | 31.9812 |  | 1
| 1 | 2 | 21 | 7.78 | 20.9 | 21.8 | 21.7 | 21.7 | 20.5 | 23.3 | 21.9 | 28.557 | 16.93 | 21 | 20.9627 | 22.3 | 27.7459 |  | 6
| 1 | 3 | 20 | 8.11 | 20 | 20.7 | 21.5 | 21.3 | 19.8 | 22.1 | 21 | 27.21 | 14.93 | 23 | 21.8245 | 21.4 | 33.5624 |  | 16
| 1 | 4 | 22 | 8.21 | 20.9 | 20.4 | 21.8 | 22.6 | 20 | 22 | 19.9 | 24.287 | 13.67 | 22 | 21.8831 | 21.6 | 30.7388 |  | 21.3278
| 1 | 5 | 20 | 8.29 | 21 | 19.4 | 20.2 | 23 | 21.3 | 23 | 21.4 | 26.516 | 13.46 | 20 | 21.1818 | 20.2 | 29.9012 |  | 123
| 1 | 6 | 21 | 7.87 | 20.2 | 21.3 | 21.1 | 22.3 | 21.6 | 23.6 | 22.3 | 26.933 | 12.73 | 20 | 21.5611 | 22.1 | 24.3953
| Mean |  | 20.6667 | 8.04 | 20.6833 | 20.45 | 21.1833 | 22.5333 | 20.7333 | 22.6167 | 21.7667 | 27.3453 | 14.5783 | 21 | 21.422 | 21.55 | 29.7208
| Median |  | 20.5 | 8.045 | 20.9 | 20.55 | 21.3 | 22.45 | 20.85 | 22.55 | 21.65 | 27.0715 | 14.3 | 20.5 | 21.3714 | 21.65 | 30.32
| Std Dev. |  | 0.8165 | 0.19819 | 0.46224 | 1.0521 | 0.61128 | 1.0596 | 0.74207 | 0.78337 | 1.4109 | 2.1019 | 1.5823 | 1.2649 | 0.38845 | 0.73959 | 3.2618
| Rel.Std.Dev. |  | 3.951% | 2.465% | 2.235% | 5.145% | 2.886% | 4.702% | 3.579% | 3.464% | 6.482% | 7.686% | 10.85% | 6.023% | 1.813% | 3.432% | 10.97%
| PDM3 |  | -3.1% | -62.3% | -3.022% | -4.116% | -0.6772% | 5.653% | -2.787% | 6.043% | 2.058% | 28.21% | -31.65% | -1.537% | 0.4419% | 1.042% | 39.35%
| Z-Score (Absolute) |  | 0.6 | 15.19 | 0.58 | 0.85 | 0 | 1.56 | 0.52 | 1.66 | 0.67 | 7.12 | 7.64 | 0.21 | 0.28 | 0.42 | 9.87
| Analytical results for Y in OREAS 143 (Certified Value 13.8 ppm) | [17 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 18 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 11.4 | 13.06 | 12.25 | 13 | 16.5 | 13.1 | 14.2 | 15.1 | 15.4 | 14.5 | 12.4 | 12.3835 | 15.2 | 13.7978 | 13.4 | 12.9115 | 1
| 1 | 2 | 12.3 | 12.72 | 11.95 | 13.2 | 16.2 | 13.4 | 14.2 | 16.4 | 16.6 | 14.7 | 12.323 | 12.1302 | 15.3 | 14.3292 | 12.4 | 13.1212 | 7
| 1 | 3 | 12.8 | 12.97 | 12.22 | 13.1 | 17 | 14 | 13.8 | 15.1 | 15.8 | 14.2 | 12.613 | 12.1927 | 16.3 | 14.592 | 12.4 | 12.884 | 16
| 1 | 4 | 12.2 | 12.95 | 11.27 | 12.6 | 16.6 | 14.1 | 14 | 14.1 | 15.8 | 14.5 | 12.6 | 12.2313 | 15.5 | 14.5388 | 12.9 | 12.3682 | 13.7663
| 1 | 5 | 13.6 | 12.89 | 12.64 | 12.7 | 15.7 | 13 | 14.6 | 14.2 | 15.4 | 14.6 | 12.851 | 12.201 | 15.3 | 14.1331 | 12.4 | 12.7365 | 124
| 1 | 6 | 12.3 | 12.97 | 12.7 | 12.8 | 16.2 | 13.4 | 14.6 | 14.5 | 16 | 15.1 | 12.923 | 12.7225 | 15 | 14.5456 | 13 | 13.7443
| Mean |  | 12.4333 | 12.9267 | 12.1717 | 12.9 | 16.3667 | 13.5 | 14.2333 | 14.9 | 15.8333 | 14.6 | 12.6183 | 12.3102 | 15.4333 | 14.3228 | 12.75 | 12.9609
| Median |  | 12.3 | 12.96 | 12.235 | 12.9 | 16.35 | 13.4 | 14.2 | 14.8 | 15.8 | 14.55 | 12.6065 | 12.2162 | 15.3 | 14.434 | 12.65 | 12.8978
| Std Dev. |  | 0.72847 | 0.11501 | 0.52327 | 0.23664 | 0.44121 | 0.45607 | 0.32042 | 0.85088 | 0.44572 | 0.29665 | 0.23758 | 0.21894 | 0.45461 | 0.30998 | 0.41833 | 0.45813
| Rel.Std.Dev. |  | 5.859% | 0.8897% | 4.299% | 1.834% | 2.696% | 3.378% | 2.251% | 5.711% | 2.815% | 2.032% | 1.883% | 1.779% | 2.946% | 2.164% | 3.281% | 3.535%
| PDM3 |  | -9.683% | -6.099% | -11.58% | -6.293% | 18.89% | -1.934% | 3.393% | 8.235% | 15.02% | 6.056% | -8.339% | -10.58% | 12.11% | 4.042% | -7.382% | -5.85%
| Z-Score (Absolute) |  | 0.56 | 0.21 | 0.74 | 0.23 | 2.2 | 0.19 | 0.7 | 1.17 | 1.83 | 0.96 | 0.43 | 0.65 | 1.54 | 0.77 | 0.34 | 0.19
| Analytical results for Yb in OREAS 143 (Certified Value 1.27 ppm) | [17 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 12 | 13 | 14 | 18 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [8 empty] | 2
3
| 1 | 1 | 1.3 | 1.63 | 1.32 | 1.6 | 1.1284 | 1.2 | 1.49 | 1.1509 | [8 empty] | 1
| 1 | 2 | 1.2 | 1.61 | 1.31 | 1.6 | 1.1619 | 1.3 | 1.4692 | 1.1442 | [8 empty] | #N/A
| 1 | 3 | 1.4 | 1.67 | 1.24 | 1.8 | 1.2154 | 1.4 | 1.4197 | 1.1828 | [8 empty] | 16
| 1 | 4 | 1.4 | 1.67 | 1.24 | 1.7 | 1.1778 | 1.2 | 1.4914 | 1.1693 | [8 empty] | 1.2672
| 1 | 5 | 1.1 | 1.61 | 1.26 | 1.7 | 1.1315 | 1.2 | 1.3874 | 1.2144 | [8 empty] | 125
| 1 | 6 | 1.2 | 1.67 | 1.28 | 1.6 | 1.1643 | 1.2 | 1.4661 | 1.3047
| Mean |  | 1.2667 | 1.6433 | 1.275 | 1.6667 | 1.1632 | 1.25 | 1.454 | 1.1944
| Median |  | 1.25 | 1.65 | 1.27 | 1.65 | 1.1631 | 1.2 | 1.4677 | 1.1761
| Std Dev. |  | 0.12111 | 0.030111 | 0.034496 | 0.08165 | 0.032143 | 0.083666 | 0.041692 | 0.05957
| Rel.Std.Dev. |  | 9.561% | 1.832% | 2.706% | 4.899% | 2.763% | 6.693% | 2.867% | 4.988%
| PDM3 |  | -0.04252% | 29.68% | 0.6151% | 31.52% | -8.207% | -1.358% | 14.74% | -5.747%
| Z-Score (Absolute) |  | 0.03 | 2.73 | 0.03 | 2.9 | 0.79 | 0.15 | 1.34 | 0.56
| Analytical results for Zn in OREAS 143 (Certified Value 100 ppm) | [17 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 0
1
| 1 | 1 | 113 | 87 | 97 | 96 | 101 | 103 | 101 | 107 | 103 | 99.48 | 94.0555 | 96 | 99.5487 | 98 | 141 | 99.89 | 1
| 1 | 2 | 112 | 85 | 93 | 97 | 100 | 102 | 102 | 107 | 109 | 96.395 | 93.856 | 102 | 97.7286 | 95 | 143 | 99.76 | #N/A
| 1 | 3 | 111 | 87 | 96 | 94 | 101 | 105 | 99 | 108 | 100 | 97.594 | 93.73 | 106 | 98.9319 | 99 | 145 | 95.64 | 16
| 1 | 4 | 127 | 87 | 95 | 97 | 102 | 101 | 104 | 110 | 102 | 98.583 | 92.6275 | 110 | 100.1266 | 100 | 145 | 103.01 | 99.8651
| 1 | 5 | 269 | 84 | 94 | 93 | 107 | 104 | 99 | 126 | 102 | 96.474 | 94.423 | 100 | 99.7138 | 91 | 142 | 97.58 | 126
| 1 | 6 | 167 | 89 | 96 | 95 | 102 | 103 | 100 | 108 | 107 | 97.227 | 95.309 | 97 | 102.9009 | 99 | 145 | 94.69
| Mean |  | 149.8333 | 86.5 | 95.1667 | 95.3333 | 102.1667 | 103 | 100.8333 | 111 | 103.8333 | 97.6255 | 94.0002 | 101.8333 | 99.8251 | 97 | 143.5 | 98.4283
| Median |  | 120 | 87 | 95.5 | 95.5 | 101.5 | 103 | 100.5 | 108 | 102.5 | 97.4105 | 93.9557 | 101 | 99.6312 | 98.5 | 144 | 98.67
| Std Dev. |  | 62.1496 | 1.7607 | 1.472 | 1.633 | 2.4833 | 1.4142 | 1.9408 | 7.4297 | 3.4303 | 1.2129 | 0.8801 | 5.3821 | 1.7224 | 3.4059 | 1.7607 | 3.078
| Rel.Std.Dev. |  | 41.48% | 2.035% | 1.547% | 1.713% | 2.431% | 1.373% | 1.925% | 6.693% | 3.304% | 1.242% | 0.9363% | 5.285% | 1.725% | 3.511% | 1.227% | 3.127%
| PDM3 |  | 50.04% | -13.38% | -4.705% | -4.538% | 2.305% | 3.139% | 0.9696% | 11.15% | 3.974% | -2.243% | -5.873% | 1.971% | -0.04002% | -2.869% | 43.69% | -1.439%
| Z-Score (Absolute) |  | 9.77 | 2.73 | 1.02 | 0.99 | 0.36 | 0.53 | 0.1 | 2.11 | 0.69 | 0.53 | 1.25 | 0.3 | 0.1 | 0.66 | 8.52 | 0.38
| Analytical results for Zr in OREAS 143 (Certified Value 86 ppm) | [17 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 16 | 18 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  |  | 0
1
| 1 | 1 | 74 | 80.3 | 82.4 | 100 | 85.1 | 90.7 | 92.6 | 94 | 87.2 | 77.398 | 85.5 | 85.5479 | 91.6 | 74.86 |  |  | 1
| 1 | 2 | 70 | 80.12 | 82.6 | 98 | 88.6 | 92.4 | 92.4 | 97.9 | 86.7 | 77.468 | 88.8 | 87.7332 | 85.9 | 73.49 |  |  | 18
| 1 | 3 | 81 | 80.19 | 81.3 | 104 | 90.6 | 93.8 | 92.1 | 93.1 | 86 | 76.937 | 90.6 | 88.5818 | 82.2 | 70.43 |  |  | 16
| 1 | 4 | 68 | 80.05 | 83.5 | 103 | 89.1 | 89.1 | 86.8 | 97.1 | 87.7 | 78.353 | 89.7 | 89.3401 | 76.5 | 74.71 |  |  | 86.0146
| 1 | 5 | 70 | 80.26 | 81.3 | 99 | 81.7 | 94.4 | 89.2 | 97 | 87.9 | 77.207 | 88.1 | 87.3984 | 83.8 | 74.91 |  |  | 127
| 1 | 6 | 80 | 80.98 | 83.6 | 101 | 84.9 | 93.9 | 92.2 | 92.7 | 90.1 | 77.835 | 88 | 88.7963 | 92.1 | 79.83
| Mean |  | 73.8333 | 80.3167 | 82.45 | 100.8333 | 86.6667 | 92.3833 | 90.8833 | 95.3 | 87.6 | 77.533 | 88.45 | 87.8996 | 85.35 | 74.705
| Median |  | 72 | 80.225 | 82.5 | 100.5 | 86.85 | 93.1 | 92.15 | 95.5 | 87.45 | 77.433 | 88.45 | 88.1575 | 84.85 | 74.785
| Std Dev. |  | 5.5287 | 0.33744 | 1.0095 | 2.3166 | 3.3279 | 2.0971 | 2.3651 | 2.2882 | 1.4057 | 0.49951 | 1.7513 | 1.353 | 5.9258 | 3.0359
| Rel.Std.Dev. |  | 7.488% | 0.4201% | 1.224% | 2.297% | 3.84% | 2.27% | 2.602% | 2.401% | 1.605% | 0.6443% | 1.98% | 1.539% | 6.943% | 4.064%
| PDM3 |  | -14.16% | -6.624% | -4.144% | 17.23% | 0.7581% | 7.404% | 5.66% | 10.8% | 1.843% | -9.861% | 2.831% | 2.191% | -0.7726% | -13.15%
| Z-Score (Absolute) |  | 1.81 | 0.93 | 0.64 | 1.86 | 0.06 | 0.71 | 0.51 | 1.11 | 0.06 | 1.3 | 0.18 | 0.1 | 0.24 | 1.69
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C905: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C942: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C978: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1015: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1033: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1051: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1069: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1087: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1105: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1123: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for S in OREAS 143 (Indicative Value 1.23 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 02 | wt.%
|  |  | IRC | 2
2
| 1 | 1 | 1.24 | 1
| 1 | 2 | 1.23 | 24
| 1 | 3 | 1.24 | 16
| 1 | 4 | 1.22 | 1.23
| 1 | 5 | 1.23 | 30
| 1 | 6 | 1.22
| Mean |  | 1.23
| Median |  | 1.23
| Std Dev. |  | 0.0089443
| Rel.Std.Dev. |  | 0.7272%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.