## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 196
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 3.5554 | 3.4849 | 3.6258 | 3.5235 | 3.5873
CaO, Calcium oxide (wt.%) | 0.3114 | 0.30091 | 0.32188 | 0.30502 | 0.31778
Co, Cobalt (ppm) | 602.05 | 589.1312 | 614.9688 | 590.1309 | 613.9691
Cr2O3, Chromium(III) oxide (wt.%) | 1.3484 | 1.3296 | 1.3672 | 1.3342 | 1.3626
Cu, Copper (ppm) | < 50 | IND | IND | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 27.56 | 27.3486 | 27.7714 | 27.399 | 27.721
K2O, Potassium oxide (wt.%) | 0.03834 | 0.031404 | 0.045277 | 0.034332 | 0.042349
MgO, Magnesium oxide (wt.%) | 12.6716 | 12.5453 | 12.7979 | 12.5952 | 12.748
MnO, Manganese oxide (wt.%) | 0.49423 | 0.48314 | 0.50532 | 0.48814 | 0.50033
Ni, Nickel (wt.%) | 2.0858 | 2.0581 | 2.1135 | 2.0699 | 2.1017
P2O5, Phosphorus(V) oxide (wt.%) | 0.0083227 | 0.0042448 | 0.012401 | IND | IND
SiO2, Silicon dioxide (wt.%) | 41.9771 | 41.5259 | 42.4282 | 41.7896 | 42.1645
Sn, Tin (ppm) | < 50 | IND | IND | IND | IND
SO3, Sulphur trioxide (wt.%) | 0.50394 | 0.47316 | 0.53471 | 0.49187 | 0.516
TiO2, Titanium dioxide (wt.%) | 0.1388 | 0.12846 | 0.14914 | IND | IND
V2O5, Vanadium(V) oxide (ppm) | 171.3033 | 141.8792 | 200.7274 | IND | IND
Zn, Zinc (ppm) | 431.5263 | 409.3743 | 453.6784 | 420.2043 | 442.8484
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 9.5076 | 8.6065 | 10.4086 | 9.3878 | 9.6273
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 1.8317 | 1.7906 | 1.8728 | 1.7936 | 1.8699
Ba, Barium (ppm) | 34.9416 | 30.9004 | 38.9828 | 32.6456 | 37.2376
Be, Beryllium (ppm) | < 1 | IND | IND | IND | IND
Bi, Bismuth (ppm) | < 0.1 | IND | IND | IND | IND
Ca, Calcium (wt.%) | 0.21895 | 0.197 | 0.2409 | 0.20815 | 0.22974
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (ppm) | 2.8756 | 2.3468 | 3.4044 | 2.581 | 3.1702
Co, Cobalt (ppm) | 574.1102 | 550.3647 | 597.8556 | 563.9549 | 584.2654
Cr, Chromium (wt.%) | 0.90003 | 0.87328 | 0.92678 | 0.87675 | 0.92331
Cs, Caesium (ppm) | 0.18819 | 0.10058 | 0.27579 | IND | IND
Cu, Copper (ppm) | 45.0113 | 40.24 | 49.7825 | 41.1861 | 48.8364
Dy, Dysprosium (ppm) | 1.5812 | 1.3605 | 1.8018 | 1.4214 | 1.7409
Er, Erbium (ppm) | 0.93545 | 0.79976 | 1.0711 | 0.85931 | 1.0116
Eu, Europium (ppm) | 0.513 | 0.41469 | 0.6113 | 0.4508 | 0.57519
Fe, Iron (wt.%) | 18.7429 | 18.3239 | 19.162 | 18.4029 | 19.083
Ga, Gallium (ppm) | 4.2168 | 2.7656 | 5.668 | 3.7965 | 4.6371
Gd, Gadolinium (ppm) | 1.7288 | 1.481 | 1.9766 | 1.5857 | 1.8719
Ge, Germanium (ppm) | 2.4027 | 1.5783 | 3.2271 | IND | IND
Ho, Holmium (ppm) | 0.31643 | 0.26948 | 0.36338 | 0.2738 | 0.35906
In, Indium (ppm) | < 0.1 | IND | IND | IND | IND
La, Lanthanum (ppm) | 10.3653 | 9.5523 | 11.1783 | 9.8932 | 10.8374
Lu, Lutetium (ppm) | 0.12231 | 0.077236 | 0.16739 | IND | IND
Mg, Magnesium (wt.%) | 7.3928 | 7.2021 | 7.5835 | 7.2613 | 7.5243
Mn, Manganese (wt.%) | 0.36777 | 0.35623 | 0.37931 | 0.35927 | 0.37627
Nd, Neodymium (ppm) | 9.2473 | 8.188 | 10.3067 | 8.7599 | 9.7348
Ni, Nickel (wt.%) | 2.0197 | 1.9372 | 2.1023 | 1.9862 | 2.0533
Pr, Praseodymium (ppm) | 2.387 | 2.1792 | 2.5949 | 2.2383 | 2.5358
Rb, Rubidium (ppm) | 1.6908 | 1.2901 | 2.0915 | IND | IND
Re, Rhenium (ppm) | < 0.1 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.17842 | 0.16079 | 0.19605 | 0.16684 | 0.19001
Sc, Scandium (ppm) | 22.0178 | 20.2889 | 23.7466 | 20.8183 | 23.2173
Se, Selenium (ppm) | < 20 | IND | IND | IND | IND
Si, Silicon (wt.%) | 19.2205 | 18.6791 | 19.7619 | 18.7495 | 19.6914
Sm, Samarium (ppm) | 1.6303 | 1.3673 | 1.8932 | IND | IND
Sr, Strontium (ppm) | 19.6984 | 16.5174 | 22.8794 | 17.9878 | 21.4091
Tb, Terbium (ppm) | 0.25437 | 0.21089 | 0.29785 | 0.21754 | 0.2912
Te, Tellurium (ppm) | < 1 | IND | IND | IND | IND
Th, Thorium (ppm) | 0.18778 | 0.16388 | 0.21167 | IND | IND
Ti, Titanium (wt.%) | 0.080561 | 0.074545 | 0.086577 | 0.076386 | 0.084736
Tl, Thallium (ppm) | < 0.5 | IND | IND | IND | IND
Tm, Thulium (ppm) | 0.13292 | 0.093293 | 0.17254 | IND | IND
V, Vanadium (ppm) | 93.5686 | 85.3163 | 101.8209 | 89.9793 | 97.158
W, Tungsten (ppm) | 2.2119 | 1.624 | 2.7997 | IND | IND
Y, Yttrium (ppm) | 8.6392 | 7.9802 | 9.2981 | 8.2995 | 8.9788
Yb, Ytterbium (ppm) | 0.8432 | 0.66996 | 1.0164 | IND | IND
Zn, Zinc (ppm) | 411.2287 | 385.4128 | 437.0445 | 400.4209 | 422.0365
Infrared Combustion
C, Carbon (wt.%) | 0.56269 | 0.54332 | 0.58206 | 0.54964 | 0.57575
S, Sulphur (wt.%) | 0.18228 | 0.17081 | 0.19376 | 0.17351 | 0.19106
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 196
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
As | ppm | < 100 | Hg | ppm | < 100 | Se | ppm | < 100
BaO | ppm | 89.0169 | In | ppm | < 100 | SrO | ppm | < 120
Bi | ppm | < 100 | La | ppm | 67.5023 | Ta | ppm | < 100
Cd | ppm | < 100 | Mo | ppm | < 50 | Te | ppm | < 100
Ce | ppm | < 80 | Na2O | wt.% | 0.11496 | Tl | ppm | < 100
Cl | ppm | 82.866 | Nb | ppm | < 50 | W | ppm | < 10
Cs | ppm | < 100 | Pb | ppm | < 50 | Y | ppm | < 39
Ga | ppm | < 100 | Rb | ppm | < 50 | Zr | ppm | 57.5833
Ge | ppm | < 100 | Sb | ppm | < 50
HfO2 | ppm | < 100 | Sc | ppm | 20.8333
Thermogravimetry
H2O- | wt.% | 5.0867
Borate / Peroxide Fusion ICP
Ag | ppm | < 0.5 | Li | ppm | 3.065 | Sb | ppm | 0.64583
As | ppm | < 100 | Mo | ppm | 1.1237 | Sn | ppm | 0.9775
B | ppm | 40.2074 | Na | wt.% | 0.13917 | Ta | ppm | 0.51773
Hf | ppm | < 1 | Nb | ppm | < 2 | U | ppm | 0.18386
Hg | ppm | < 5 | P | wt.% | 0.0070002 | Zr | ppm | 9.5192
K | wt.% | 0.09114 | Pb | ppm | 30.7
4-Acid Digestion
Co | ppm | 589.1552 | Mg | wt.% | 7.2439
Fe | wt.% | 17.8223 | Ni | wt.% | 2.0473
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 196
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 3.5554 | 0.076553 | 3.4023 | 3.7085 | 3.3257 | 3.785 | 2.153% | 4.306% | 6.459% | 3.3776 | 3.7331
CaO, wt.% | 0.3114 | 0.0063721 | 0.29865 | 0.32414 | 0.29228 | 0.33051 | 2.046% | 4.093% | 6.139% | 0.29583 | 0.32697
Co, ppm | 602.05 | 9.4468 | 583.1565 | 620.9436 | 573.7097 | 630.3904 | 1.569% | 3.138% | 4.707% | 571.9475 | 632.1525
Cr2O3, wt.% | 1.3484 | 0.023898 | 1.3006 | 1.3962 | 1.2767 | 1.4201 | 1.772% | 3.545% | 5.317% | 1.281 | 1.4158
Cu, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Fe2O3, wt.% | 27.56 | 0.19045 | 27.1791 | 27.9409 | 26.9887 | 28.1314 | 0.691% | 1.382% | 2.073% | 26.182 | 28.938
K2O, wt.% | 0.03834 | 0.0052037 | 0.027933 | 0.048748 | 0.022729 | 0.053951 | 13.57% | 27.14% | 40.72% | 0.036423 | 0.040257
MgO, wt.% | 12.6716 | 0.15846 | 12.3547 | 12.9885 | 12.1962 | 13.147 | 1.25% | 2.501% | 3.751% | 12.038 | 13.3052
MnO, wt.% | 0.49423 | 0.010801 | 0.47263 | 0.51583 | 0.46183 | 0.52664 | 2.185% | 4.371% | 6.556% | 0.46952 | 0.51894
Ni, wt.% | 2.0858 | 0.040738 | 2.0043 | 2.1673 | 1.9636 | 2.208 | 1.953% | 3.906% | 5.859% | 1.9815 | 2.1901
P2O5, wt.% | 0.0083227 | 0.0025313 | 0.00326 | 0.013385 | 0.00072873 | 0.015917 | 30.41% | 60.83% | 91.24% | 0.0079065 | 0.0087388
SiO2, wt.% | 41.9771 | 0.69534 | 40.5864 | 43.3677 | 39.8911 | 44.0631 | 1.656% | 3.313% | 4.969% | 39.8782 | 44.0759
Sn, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SO3, wt.% | 0.50394 | 0.041969 | 0.42 | 0.58788 | 0.37803 | 0.62985 | 8.328% | 16.66% | 24.98% | 0.47874 | 0.52913
TiO2, wt.% | 0.1388 | 0.010704 | 0.11739 | 0.16021 | 0.10669 | 0.17091 | 7.712% | 15.42% | 23.13% | 0.13186 | 0.14574
V2O5, ppm | 171.3033 | 14.2861 | 142.7312 | 199.8755 | 128.4451 | 214.1616 | 8.34% | 16.68% | 25.02% | 162.7382 | 179.8685
Zn, ppm | 431.5263 | 21.8442 | 387.8379 | 475.2148 | 365.9936 | 497.059 | 5.062% | 10.12% | 15.19% | 409.95 | 453.1026
Thermogravimetry
LOI1000, wt.% | 9.5076 | 1.8159 | 5.8757 | 13.1394 | 4.0598 | 14.9553 | 19.1% | 38.2% | 57.3% | 9.0322 | 9.9829
Borate / Peroxide Fusion ICP
Al, wt.% | 1.8317 | 0.038534 | 1.7547 | 1.9088 | 1.7161 | 1.9473 | 2.104% | 4.207% | 6.311% | 1.7401 | 1.9233
Ba, ppm | 34.9416 | 3.1139 | 28.7138 | 41.1695 | 25.5999 | 44.2834 | 8.912% | 17.82% | 26.74% | 33.1946 | 36.6887
Be, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Bi, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ca, wt.% | 0.21895 | 0.030341 | 0.15826 | 0.27963 | 0.12792 | 0.30997 | 13.86% | 27.72% | 41.57% | 0.208 | 0.22989
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 2.8756 | 0.26851 | 2.3386 | 3.4126 | 2.0701 | 3.6811 | 9.338% | 18.68% | 28.01% | 2.7318 | 3.0194
Co, ppm | 574.1102 | 28.0876 | 517.935 | 630.2853 | 489.8474 | 658.3729 | 4.892% | 9.785% | 14.68% | 545.4047 | 602.8157
Cr, wt.% | 0.90003 | 0.029645 | 0.84074 | 0.95932 | 0.81109 | 0.98896 | 3.294% | 6.588% | 9.882% | 0.85503 | 0.94503
Cs, ppm | 0.18819 | 0.045532 | 0.097122 | 0.27925 | 0.05159 | 0.32478 | 24.2% | 48.39% | 72.59% | 0.17878 | 0.1976
Cu, ppm | 45.0113 | 3.2374 | 38.5364 | 51.4861 | 35.299 | 54.7235 | 7.192% | 14.38% | 21.58% | 42.7607 | 47.2618
Dy, ppm | 1.5812 | 0.20263 | 1.1759 | 1.9864 | 0.97329 | 2.189 | 12.81% | 25.63% | 38.44% | 1.5021 | 1.6602
Er, ppm | 0.93545 | 0.059845 | 0.81576 | 1.0551 | 0.75592 | 1.115 | 6.397% | 12.79% | 19.19% | 0.88868 | 0.98222
Eu, ppm | 0.513 | 0.070839 | 0.37132 | 0.65467 | 0.30048 | 0.72551 | 13.81% | 27.62% | 41.43% | 0.48735 | 0.53864
Fe, wt.% | 18.7429 | 0.34244 | 18.0581 | 19.4278 | 17.7156 | 19.7703 | 1.827% | 3.654% | 5.481% | 17.8058 | 19.6801
Ga, ppm | 4.2168 | 0.91792 | 2.3809 | 6.0526 | 1.463 | 6.9705 | 21.77% | 43.54% | 65.31% | 4.0059 | 4.4276
Gd, ppm | 1.7288 | 0.12711 | 1.4746 | 1.983 | 1.3475 | 2.1102 | 7.353% | 14.71% | 22.06% | 1.6424 | 1.8153
Ge, ppm | 2.4027 | 0.45046 | 1.5018 | 3.3036 | 1.0513 | 3.7541 | 18.75% | 37.5% | 56.24% | 2.2826 | 2.5228
Ho, ppm | 0.31643 | 0.018888 | 0.27865 | 0.3542 | 0.25976 | 0.37309 | 5.969% | 11.94% | 17.91% | 0.30061 | 0.33225
In, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
La, ppm | 10.3653 | 0.49183 | 9.3817 | 11.349 | 8.8898 | 11.8408 | 4.745% | 9.49% | 14.23% | 9.8471 | 10.8836
Lu, ppm | 0.12231 | 0.032498 | 0.057317 | 0.18731 | 0.024819 | 0.21981 | 26.57% | 53.14% | 79.71% | 0.1162 | 0.12843
Mg, wt.% | 7.3928 | 0.12034 | 7.1521 | 7.6335 | 7.0318 | 7.7538 | 1.628% | 3.256% | 4.884% | 7.0231 | 7.7624
Mn, wt.% | 0.36777 | 0.014593 | 0.33858 | 0.39696 | 0.32399 | 0.41155 | 3.968% | 7.936% | 11.9% | 0.34938 | 0.38616
Nd, ppm | 9.2473 | 0.3848 | 8.4777 | 10.0169 | 8.0929 | 10.4017 | 4.161% | 8.322% | 12.48% | 8.785 | 9.7097
Ni, wt.% | 2.0197 | 0.10588 | 1.808 | 2.2315 | 1.7021 | 2.3374 | 5.242% | 10.48% | 15.73% | 1.9187 | 2.1207
Pr, ppm | 2.387 | 0.10792 | 2.1712 | 2.6029 | 2.0633 | 2.7108 | 4.521% | 9.042% | 13.56% | 2.2677 | 2.5064
Rb, ppm | 1.6908 | 0.2015 | 1.2878 | 2.0938 | 1.0863 | 2.2953 | 11.92% | 23.84% | 35.75% | 1.6063 | 1.7753
Re, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.17842 | 0.021903 | 0.13462 | 0.22223 | 0.11272 | 0.24413 | 12.28% | 24.55% | 36.83% | 0.1695 | 0.18735
Sc, ppm | 22.0178 | 1.4413 | 19.1352 | 24.9003 | 17.694 | 26.3415 | 6.546% | 13.09% | 19.64% | 20.9169 | 23.1187
Se, ppm | < 20 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Si, wt.% | 19.2205 | 0.56786 | 18.0848 | 20.3562 | 17.5169 | 20.9241 | 2.954% | 5.909% | 8.863% | 18.2595 | 20.1815
Sm, ppm | 1.6303 | 0.15041 | 1.3294 | 1.9311 | 1.179 | 2.0815 | 9.226% | 18.45% | 27.68% | 1.5488 | 1.7118
Sr, ppm | 19.6984 | 2.1748 | 15.3488 | 24.048 | 13.174 | 26.2228 | 11.04% | 22.08% | 33.12% | 18.7135 | 20.6833
Tb, ppm | 0.25437 | 0.025384 | 0.2036 | 0.30514 | 0.17822 | 0.33053 | 9.979% | 19.96% | 29.94% | 0.24165 | 0.26709
Te, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 0.18778 | 0.027901 | 0.13198 | 0.24358 | 0.10407 | 0.27148 | 14.86% | 29.72% | 44.58% | 0.17839 | 0.19717
Ti, wt.% | 0.080561 | 0.0038497 | 0.072862 | 0.08826 | 0.069012 | 0.09211 | 4.779% | 9.557% | 14.34% | 0.076533 | 0.084589
Tl, ppm | < 0.5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tm, ppm | 0.13292 | 0.028866 | 0.075184 | 0.19065 | 0.046318 | 0.21952 | 21.72% | 43.44% | 65.15% | 0.12627 | 0.13956
V, ppm | 93.5686 | 7.7371 | 78.0944 | 109.0428 | 70.3573 | 116.7799 | 8.269% | 16.54% | 24.81% | 88.8902 | 98.247
W, ppm | 2.2119 | 0.51018 | 1.1915 | 3.2322 | 0.68133 | 3.7424 | 23.07% | 46.13% | 69.2% | 2.1013 | 2.3225
Y, ppm | 8.6392 | 0.41281 | 7.8135 | 9.4648 | 7.4007 | 9.8776 | 4.778% | 9.557% | 14.34% | 8.2072 | 9.0711
Yb, ppm | 0.8432 | 0.094458 | 0.65428 | 1.0321 | 0.55982 | 1.1266 | 11.2% | 22.4% | 33.61% | 0.80104 | 0.88536
Zn, ppm | 411.2287 | 34.1853 | 342.858 | 479.5994 | 308.6727 | 513.7847 | 8.313% | 16.63% | 24.94% | 390.6673 | 431.7901
Infrared Combustion
C, wt.% | 0.56269 | 0.022428 | 0.51784 | 0.60755 | 0.49541 | 0.62998 | 3.986% | 7.972% | 11.96% | 0.53456 | 0.59083
S, wt.% | 0.18228 | 0.011385 | 0.15952 | 0.20505 | 0.14813 | 0.21644 | 6.245% | 12.49% | 18.74% | 0.17317 | 0.1914
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 196
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PSF*XRF | pyro-sulfate fusion with X-ray fluorescence spectroscopy
Text Values:
- | Not Analysed (Lab 19)
N.A. | Not Applicable (Lab 21)
NA | Not Applicable (Lab 02)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 196
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Labwest Minerals Analysis, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
Ni Lab, Pouembout, New Caledonia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Fusion XRF
| Analytical results for Al2O3 in OREAS 196 (Certified Value 3.56 wt.%) | [20 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 18 | 20 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 3.44 | 3.63 | 3.478 | 3.5785 | 3.63 | 3.6 | 3.61 | 3.59 | 3.52 | 3.61 | 3.44 | 3.52 | 3.53 | 3.47 | 3.64 | 3.6 | 3.666 | 3.582 |  | 1
| 1 | 2 | 3.37 | 3.59 | 3.5024 | 3.5078 | 3.56 | 3.6 | 3.59 | 3.58 | 3.47 | 3.66 | 3.43 | 3.54 | 3.45 | 3.44 | 3.67 | 3.6 | 3.666 | 3.593 |  | #N/A
| 1 | 3 | 3.42 | 3.62 | 3.4951 | 3.5357 | 3.55 | 3.56 | 3.59 | 3.64 | 3.61 | 3.61 | 3.43 | 3.51 | 3.42 | 3.46 | 3.56 | 3.6 | 3.685 | 3.591 |  | 16
| 1 | 4 | 3.33 | 3.64 | 3.5018 | 3.5015 | 3.58 | 3.5 | 3.59 | 3.6 | 3.54 | 3.63 | 3.44 | 3.53 | 3.53 | 3.47 | 3.61 | 3.6 | 3.685 | 3.602 |  | 3.5554
| 1 | 5 | 3.5 | 3.64 | 3.5305 | 3.5531 | 3.61 | 3.57 | 3.62 | 3.6 | 3.52 | 3.66 | 3.41 | 3.52 | 3.52 | 3.46 | 3.62 | 3.6 | 3.685 | 3.594 |  | 7
| 1 | 6 | 3.62 | 3.66 | 3.5067 | 3.5817 | 3.43 | 3.61 | 3.61 | 3.56 | 3.53 | 3.62 | 3.44 | 3.49 | 3.49 | 3.48 | 3.55 | 3.59 | 3.685 | 3.595
| Mean |  | 3.4467 | 3.63 | 3.5024 | 3.543 | 3.56 | 3.5733 | 3.6017 | 3.595 | 3.5317 | 3.6317 | 3.4317 | 3.5183 | 3.49 | 3.4633 | 3.6083 | 3.5983 | 3.6787 | 3.5928
| Median |  | 3.43 | 3.635 | 3.5021 | 3.5444 | 3.57 | 3.585 | 3.6 | 3.595 | 3.525 | 3.625 | 3.435 | 3.52 | 3.505 | 3.465 | 3.615 | 3.6 | 3.685 | 3.5935
| Std Dev. |  | 0.10309 | 0.023664 | 0.017067 | 0.034278 | 0.070427 | 0.040825 | 0.013292 | 0.026646 | 0.04535 | 0.023166 | 0.01169 | 0.017224 | 0.046043 | 0.013663 | 0.046224 | 0.0040825 | 0.0098116 | 0.0064936
| Rel.Std.Dev. |  | 2.991% | 0.6519% | 0.4873% | 0.9675% | 1.978% | 1.142% | 0.369% | 0.7412% | 1.284% | 0.6379% | 0.3407% | 0.4896% | 1.319% | 0.3945% | 1.281% | 0.1135% | 0.2667% | 0.1807%
| PDM3 |  | -3.057% | 2.099% | -1.489% | -0.3464% | 0.1304% | 0.5055% | 1.302% | 1.115% | -0.6665% | 2.146% | -3.479% | -1.041% | -1.838% | -2.588% | 1.49% | 1.209% | 3.468% | 1.054%
| Z-Score (Absolute) |  | 1.8 | 0.95 | 0.96 | 0.35 | 0.1 | 0.1 | 0.52 | 0.42 | 0.52 | 0.97 | 2.02 | 0.72 | 1.15 | 1.55 | 0.62 | 0.47 | 1.67 | 0.39
| Analytical results for As in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 07 | 08 | 09 | 14 | 15 | [13 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 50 | < 100 | < 100 | < 100 | 20 | < 100 | [13 empty] | 1
| 1 | 2 | < 50 | < 100 | < 100 | < 100 | 10 | < 100 | [13 empty] | 1
| 1 | 3 | < 50 | < 100 | < 100 | 100 | 10 | < 100 | [13 empty] | 16
| 1 | 4 | < 50 | < 100 | < 100 | < 100 | 20 | < 100 | [13 empty] | < 100
| 1 | 5 | < 50 | < 100 | < 100 | < 100 | < 10 | < 100 | [13 empty] | 7
| 1 | 6 | < 50 | < 100 | < 100 | < 100 | 20 | < 100
| Mean |  |  |  |  | 100 | 16
| Median |  |  |  |  | 100 | 20
| Std Dev. | [5 empty] | 5.4772
| Rel.Std.Dev. | [5 empty] | 34.23%
PDM3
| Z-Score (Absolute) |  | 4.05 | 0 | 0 | 1.35 | 5.8 | 0
| Analytical results for BaO in OREAS 196 (Indicative Value 89 ppm) | [20 empty] | Indicative
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 14 | 15 | 16 | 20 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | 100 | < 100 | 100 | 145 | < 100 | < 100 | < 100 | [9 empty] | 1
| 1 | 2 | < 100 | < 50 | 50.9326 | 100 | 100 | 100 | 33 | < 100 | < 100 | < 100 | [9 empty] | 2
| 1 | 3 | < 100 | < 50 | 64.1544 | 100 | < 100 | 100 | 67 | < 100 | < 100 | < 100 | [9 empty] | 16
| 1 | 4 | < 100 | < 50 | 52.7049 | < 100 | 100 | < 100 | 123 | < 100 | < 100 | < 100 | [9 empty] | 89.0169
| 1 | 5 | < 100 | < 50 | < 50 | 100 | 100 | 100 | 134 | < 100 | < 100 | < 100 | [9 empty] | 8
| 1 | 6 | < 100 | < 50 | 69.2965 | 100 | < 100 | < 100 | 134 | < 100 | < 100 | < 100
| Mean |  |  |  | 59.2721 | 100 | 100 | 100 | 106
| Median |  |  |  | 58.4297 | 100 | 100 | 100 | 128.5
| Std Dev. |  |  |  | 8.8882 | 0 | 0 | 0 | 45.2283
| Rel.Std.Dev. |  |  |  | 15% | 0% | 0% | 0% | 42.67%
| PDM3 |  |  |  | -33.41% | 12.34% | 12.34% | 12.34% | 19.08%
| Z-Score (Absolute) |  | 0 | 1.24 | 0.11 | 2.07 | 1.24 | 1.66 | 2.78 | 0 | 0 | 0
| Analytical results for Bi in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 07 | 08 | 09 | 15 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 200 | < 100 | [15 empty] | 1
| 1 | 2 | < 100 | < 100 | 200 | < 100 | [15 empty] | 3
| 1 | 3 | < 100 | < 100 | 200 | < 100 | [15 empty] | 16
| 1 | 4 | < 100 | < 100 | 100 | < 100 | [15 empty] | < 100
| 1 | 5 | < 100 | < 100 | 200 | < 100 | [15 empty] | 9
| 1 | 6 | < 100 | < 100 | 100 | < 100
| Mean |  |  |  | 166.6667
| Median |  |  |  | 200
| Std Dev. |  |  |  | 51.6398
| Rel.Std.Dev. |  |  |  | 30.98%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for CaO in OREAS 196 (Certified Value 0.311 wt.%) | [20 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 18 | 20 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.27 | 0.31 | 0.3044 | 0.31187 | 0.32 | 0.31 | 0.32 | 0.3 | 0.282 | 0.32 | 0.31 | 0.32 | 0.37 | 0.27 | 0.27 | 0.3 | 0.308 | 0.315 |  | 1
| 1 | 2 | 0.27 | 0.32 | 0.3029 | 0.31482 | 0.31 | 0.31 | 0.32 | 0.3 | 0.279 | 0.32 | 0.3 | 0.3 | 0.37 | 0.27 | 0.27 | 0.31 | 0.308 | 0.313 |  | #N/A
| 1 | 3 | 0.28 | 0.31 | 0.30291 | 0.31301 | 0.31 | 0.31 | 0.32 | 0.3 | 0.281 | 0.32 | 0.32 | 0.31 | 0.34 | 0.27 | 0.27 | 0.31 | 0.308 | 0.313 |  | 16
| 1 | 4 | 0.27 | 0.31 | 0.30241 | 0.31073 | 0.31 | 0.3 | 0.31 | 0.3 | 0.278 | 0.32 | 0.3 | 0.32 | 0.36 | 0.27 | 0.27 | 0.31 | 0.308 | 0.315 |  | 0.3114
| 1 | 5 | 0.28 | 0.31 | 0.30749 | 0.31449 | 0.3 | 0.31 | 0.32 | 0.3 | 0.282 | 0.32 | 0.31 | 0.31 | 0.35 | 0.27 | 0.28 | 0.3 | 0.322 | 0.314 |  | 8
| 1 | 6 | 0.29 | 0.31 | 0.30694 | 0.31673 | 0.32 | 0.31 | 0.37 | 0.3 | 0.277 | 0.32 | 0.31 | 0.31 | 0.37 | 0.27 | 0.27 | 0.3 | 0.308 | 0.313
| Mean |  | 0.27667 | 0.31167 | 0.30451 | 0.31361 | 0.31167 | 0.30833 | 0.32667 | 0.3 | 0.27983 | 0.32 | 0.30833 | 0.31167 | 0.36 | 0.27 | 0.27167 | 0.305 | 0.31033 | 0.31383
| Median |  | 0.275 | 0.31 | 0.30366 | 0.31375 | 0.31 | 0.31 | 0.32 | 0.3 | 0.28 | 0.32 | 0.31 | 0.31 | 0.365 | 0.27 | 0.27 | 0.305 | 0.308 | 0.3135
| Std Dev. |  | 0.008165 | 0.0040825 | 0.0022056 | 0.0021749 | 0.0075277 | 0.0040825 | 0.021602 | 0 | 0.002137 | 0 | 0.0075277 | 0.0075277 | 0.012649 | 0 | 0.0040825 | 0.0054772 | 0.0057155 | 0.00098319
| Rel.Std.Dev. |  | 2.951% | 1.31% | 0.7243% | 0.6935% | 2.415% | 1.324% | 6.613% | 0% | 0.7637% | 0% | 2.441% | 2.415% | 3.514% | 0% | 1.503% | 1.796% | 1.842% | 0.3133%
| PDM3 |  | -11.15% | 0.08636% | -2.212% | 0.7094% | 0.08636% | -0.9841% | 4.903% | -3.66% | -10.14% | 2.762% | -0.9841% | 0.08636% | 15.61% | -13.29% | -12.76% | -2.055% | -0.3418% | 0.7821%
| Z-Score (Absolute) |  | 4.38 | 0.2 | 0.74 | 0.45 | 0.2 | 0.24 | 2.16 | NA | 3.97 | 1.29 | 0.24 | 0.2 | 6.52 | 5.25 | 5.03 | 0.67 | 0 | 0.48
| Analytical results for Cd in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 4
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 10
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 196 (Indicative Value < 80 ppm) | [20 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 07 | 08 | 09 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 80 | < 80 | 80 | [16 empty] | 1
| 1 | 2 | < 80 | < 80 | < 80 | [16 empty] | 5
| 1 | 3 | < 80 | < 80 | 80 | [16 empty] | 16
| 1 | 4 | < 80 | < 80 | < 80 | [16 empty] | < 80
| 1 | 5 | < 80 | < 80 | 80 | [16 empty] | 11
| 1 | 6 | < 80 | < 80 | < 80
| Mean |  |  |  | 80
| Median |  |  |  | 80
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Cl in OREAS 196 (Indicative Value 83 ppm) | [20 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 04 | 14 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | [17 empty] | 0
0
| 1 | 1 | 53.2673 | 130 | [17 empty] | 1
| 1 | 2 | 72.6483 | 130 | [17 empty] | 6
| 1 | 3 | 74.6718 | 100 | [17 empty] | 16
| 1 | 4 | < 50 | 150 | [17 empty] | 82.866
| 1 | 5 | < 50 | 90 | [17 empty] | 12
| 1 | 6 | 63.8045 | 80
| Mean |  | 66.098 | 113.3333
| Median |  | 68.2264 | 115
| Std Dev. |  | 9.7691 | 27.3252
| Rel.Std.Dev. |  | 14.78% | 24.11%
| PDM3 |  | -20.24% | 36.77%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 196 (Certified Value 602 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 18 | 20 | 25 | 26 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 680 | 600 | 596.5934 | 580 | 600 | 600 | 600 | 610 | 600 | 610 | 600 | 600 | 500 | 600 | 600 | 600 |  |  |  | 1
| 1 | 2 | 690 | 600 | 601.0047 | 580 | 600 | 600 | 600 | 610 | 600 | 609 | 610 | 600 | 500 | 600 | 600 | 590 |  |  |  | 21
| 1 | 3 | 680 | 610 | 609.4666 | 595 | 600 | 600 | 600 | 620 | 600 | 608 | 610 | 600 | 500 | 600 | 600 | 590 |  |  |  | 16
| 1 | 4 | 690 | 600 | 600.8363 | 595 | 600 | 600 | 600 | 620 | 600 | 601 | 590 | 600 | 500 | 600 | 610 | 600 |  |  |  | 602.05
| 1 | 5 | 700 | 600 | 612.4149 | 580 | 600 | 600 | 600 | 610 | 500 | 608 | 610 | 600 | 500 | 600 | 610 | 600 |  |  |  | 9
| 1 | 6 | 670 | 610 | 606.1427 | 587 | 600 | 600 | 600 | 610 | 600 | 609 | 590 | 600 | 500 | 600 | 600 | 600
| Mean |  | 685 | 603.3333 | 604.4098 | 586.1667 | 600 | 600 | 600 | 613.3333 | 583.3333 | 607.5 | 601.6667 | 600 | 500 | 600 | 603.3333 | 596.6667
| Median |  | 685 | 600 | 603.5737 | 583.5 | 600 | 600 | 600 | 610 | 600 | 608.5 | 605 | 600 | 500 | 600 | 600 | 600
| Std Dev. |  | 10.4881 | 5.164 | 5.9685 | 7.3598 | 0 | 0 | 0 | 5.164 | 40.8248 | 3.2711 | 9.8319 | 0 | 0 | 0 | 5.164 | 5.164
| Rel.Std.Dev. |  | 1.531% | 0.8559% | 0.9875% | 1.256% | 0% | 0% | 0% | 0.842% | 6.999% | 0.5385% | 1.634% | 0% | 0% | 0% | 0.8559% | 0.8655%
| PDM3 |  | 13.78% | 0.2132% | 0.3919% | -2.638% | -0.3405% | -0.3405% | -0.3405% | 1.874% | -3.109% | 0.9052% | -0.06368% | -0.3405% | -16.95% | -0.3405% | 0.2132% | -0.8942%
| Z-Score (Absolute) |  | 13.22 | 0 | 0.17 | 2.78 | NA | NA | NA | 1.62 | NA | 0.67 | 0.27 | NA | NA | NA | 0 | 1.08
Results from laboratories 7, 8, 9, 12, 15, 18 and 20 were removed due to their 100 ppm reading resolution.
| Analytical results for Cr2O3 in OREAS 196 (Certified Value 1.35 wt.%) | [20 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.325 | 1.35 | 1.3517 | 1.3678 | 1.28 | 1.33 | 1.33 | 1.359 | 1.342 | 1.349 | 1.33 | 1.3987 | 1.3913 | 1.38 | 1.31 | 1.34 | 1.35 | 1.3709 | 1.343 | 1
| 1 | 2 | 1.31 | 1.356 | 1.3517 | 1.3473 | 1.26 | 1.33 | 1.359 | 1.345 | 1.33 | 1.346 | 1.34 | 1.3449 | 1.3913 | 1.37 | 1.31 | 1.32 | 1.35 | 1.3767 | 1.34 | #N/A
| 1 | 3 | 1.307 | 1.36 | 1.3492 | 1.3697 | 1.27 | 1.345 | 1.359 | 1.374 | 1.322 | 1.364 | 1.33 | 1.3809 | 1.3884 | 1.31 | 1.3 | 1.34 | 1.35 | 1.3709 | 1.336 | 16
| 1 | 4 | 1.29 | 1.351 | 1.3507 | 1.3651 | 1.27 | 1.315 | 1.359 | 1.359 | 1.353 | 1.367 | 1.34 | 1.3259 | 1.3767 | 1.33 | 1.3 | 1.34 | 1.36 | 1.3782 | 1.344 | 1.3484
| 1 | 5 | 1.346 | 1.368 | 1.3587 | 1.3683 | 1.26 | 1.345 | 1.359 | 1.374 | 1.337 | 1.358 | 1.33 | 1.3692 | 1.3855 | 1.31 | 1.31 | 1.34 | 1.35 | 1.3723 | 1.345 | 10
| 1 | 6 | 1.394 | 1.358 | 1.3522 | 1.3664 | 1.27 | 1.33 | 1.359 | 1.345 | 1.361 | 1.359 | 1.33 | 1.3759 | 1.3943 | 1.28 | 1.3 | 1.34 | 1.35 | 1.365 | 1.343
| Mean |  | 1.3287 | 1.3572 | 1.3524 | 1.3641 | 1.2683 | 1.3325 | 1.3542 | 1.3593 | 1.3408 | 1.3572 | 1.3333 | 1.3659 | 1.3879 | 1.33 | 1.305 | 1.3367 | 1.3517 | 1.3723 | 1.3418
| Median |  | 1.3175 | 1.357 | 1.3517 | 1.3671 | 1.27 | 1.33 | 1.359 | 1.359 | 1.3395 | 1.3585 | 1.33 | 1.3725 | 1.3899 | 1.32 | 1.305 | 1.34 | 1.35 | 1.3716 | 1.343
| Std Dev. |  | 0.037136 | 0.0065853 | 0.0032735 | 0.0083625 | 0.0075277 | 0.011292 | 0.011839 | 0.012972 | 0.014442 | 0.0082321 | 0.005164 | 0.026278 | 0.0062535 | 0.038471 | 0.0054772 | 0.008165 | 0.0040825 | 0.0047196 | 0.0033116
| Rel.Std.Dev. |  | 2.795% | 0.4852% | 0.2421% | 0.613% | 0.5935% | 0.8474% | 0.8743% | 0.9543% | 1.077% | 0.6066% | 0.3873% | 1.924% | 0.4506% | 2.893% | 0.4197% | 0.6108% | 0.302% | 0.3439% | 0.2468%
| PDM3 |  | -1.462% | 0.6512% | 0.2953% | 1.167% | -5.937% | -1.178% | 0.4288% | 0.8119% | -0.5601% | 0.6512% | -1.116% | 1.3% | 2.932% | -1.364% | -3.218% | -0.8691% | 0.2433% | 1.776% | -0.4859%
| Z-Score (Absolute) |  | 1.09 | 0.26 | 0.03 | 0.59 | 3.94 | 0.91 | 0.13 | 0.36 | 0.51 | 0.26 | 0.87 | 0.67 | 1.72 | 1.03 | 2.21 | 0.71 | 0 | 0.98 | 0.47
| Analytical results for Cs in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Cs | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 7
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 196 (Certified Value < 50 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 18 | 20 | 26 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 160 | < 50 | < 40 | < 50 | < 50 | < 50 |  | 3 | 30 | < 100 | 100 | < 80 | 32.7528 | [6 empty] | 1
| 1 | 2 | < 160 | < 50 | < 40 | < 50 | 50 | < 50 |  |  | 90 | < 100 | 100 | < 80 | 31.954 | [6 empty] | #N/A
| 1 | 3 | < 160 | < 50 | 40 | < 50 | < 50 | < 50 |  | 24 | 20 | < 100 | 100 | < 80 | 32.7528 | [6 empty] | 16
| 1 | 4 | < 160 | < 50 | < 40 | 90 | < 50 | < 50 |  | 22 | < 10 | < 100 | 100 | < 80 | 32.7528 | [6 empty] | < 50
| 1 | 5 | < 160 | < 50 | < 40 | < 50 | < 50 | < 50 | 50 | 76 | 20 | < 100 | 100 | < 80 | 31.954 | [6 empty] | 11
| 1 | 6 | < 160 | < 50 | < 40 | < 50 | < 50 | < 50 |  |  | < 10 | < 100 | 100 | < 80 | 32.7528
| Mean |  |  |  | 40 | 90 | 50 |  | 50 | 31.25 | 40 |  | 100 |  | 32.4866
| Median |  |  |  | 40 | 90 | 50 |  | 50 | 23 | 25 |  | 100 |  | 32.7528
| Std Dev. | [8 empty] | 31.2983 | 33.665 |  | 0 |  | 0.41252
| Rel.Std.Dev. | [8 empty] | 100.2% | 84.16% |  | 0% |  | 1.27%
PDM3
| Z-Score (Absolute) |  | 4.71 | 0.67 | 0.84 | 0.39 | 0.27 | 0.67 | 1.78 | 0.06 | 0.35 | 1.78 | NA | 0.79 | 0.06
Results from laboratory 18 were removed due to their 100 ppm reading resolution.
| Analytical results for Fe2O3 in OREAS 196 (Certified Value 27.56 wt.%) | [20 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 27.6 | 27.65 | 27.3183 | 27.4709 | 26.89 | 27.522 | 27.593 | 27.593 | 27.536 | 27.74 | 26.77 | 27.8713 | 27.65 | 26.52 | 26.44 | 28.2 | 27.48 | 27.593 | 27.7548 | 1
| 1 | 2 | 27.4 | 28 | 27.3762 | 27.2938 | 26.48 | 27.522 | 27.522 | 27.522 | 27.596 | 27.7 | 26.73 | 27.3856 | 27.622 | 26.84 | 26.35 | 28.16 | 27.46 | 27.736 | 27.6475 | #N/A
| 1 | 3 | 27.57 | 27.77 | 27.3358 | 27.465 | 26.93 | 27.593 | 27.45 | 27.736 | 27.532 | 27.85 | 26.73 | 27.8436 | 27.622 | 26.6 | 26.33 | 28.15 | 27.37 | 27.45 | 27.6804 | 16
| 1 | 4 | 27.08 | 27.76 | 27.3923 | 27.4672 | 26.7 | 27.379 | 27.379 | 27.45 | 27.546 | 27.86 | 26.87 | 26.8339 | 27.45 | 26.86 | 26.22 | 28.3 | 27.45 | 27.736 | 27.7991 | 27.56
| 1 | 5 | 27.07 | 27.68 | 27.6063 | 27.4067 | 26.54 | 27.593 | 27.593 | 27.45 | 27.528 | 27.81 | 26.62 | 27.877 | 27.593 | 26.53 | 26.36 | 28.18 | 27.43 | 27.736 | 27.7576 | 12
| 1 | 6 | 27.03 | 27.81 | 27.4193 | 27.3131 | 26.98 | 27.45 | 27.45 | 27.236 | 27.553 | 27.89 | 26.8 | 27.6889 | 27.393 | 26.72 | 26.3 | 28.14 | 27.42 | 27.593 | 27.7562
| Mean |  | 27.2917 | 27.7783 | 27.408 | 27.4028 | 26.7533 | 27.5098 | 27.4978 | 27.4978 | 27.5485 | 27.8083 | 26.7533 | 27.5834 | 27.555 | 26.6783 | 26.3333 | 28.1883 | 27.435 | 27.6407 | 27.7326
| Median |  | 27.24 | 27.765 | 27.3842 | 27.4359 | 26.795 | 27.522 | 27.486 | 27.486 | 27.541 | 27.83 | 26.75 | 27.7662 | 27.6075 | 26.66 | 26.34 | 28.17 | 27.44 | 27.6645 | 27.7555
| Std Dev. |  | 0.26332 | 0.12384 | 0.10388 | 0.080758 | 0.21182 | 0.083459 | 0.086481 | 0.16703 | 0.025026 | 0.07414 | 0.083586 | 0.41201 | 0.1065 | 0.15105 | 0.072572 | 0.058793 | 0.038341 | 0.11676 | 0.056664
| Rel.Std.Dev. |  | 0.9648% | 0.4458% | 0.379% | 0.2947% | 0.7917% | 0.3034% | 0.3145% | 0.6074% | 0.09084% | 0.2666% | 0.3124% | 1.494% | 0.3865% | 0.5662% | 0.2756% | 0.2086% | 0.1398% | 0.4224% | 0.2043%
| PDM3 |  | -0.9737% | 0.7922% | -0.5514% | -0.5704% | -2.927% | -0.182% | -0.2256% | -0.2256% | -0.04175% | 0.901% | -2.927% | 0.08478% | -0.01816% | -3.199% | -4.451% | 2.28% | -0.4536% | 0.2927% | 0.6263%
| Z-Score (Absolute) |  | 0.97 | 1.32 | 0.42 | 0.45 | 3.51 | 0.06 | 0 | 0 | 0.24 | 1.46 | 3.51 | 0.4 | 0.27 | 3.87 | 5.49 | 3.26 | 0.3 | 0.67 | 1.11
| Analytical results for Ga in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ga | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 8
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 9
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 15
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 03 | 07 | 08 | 09 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [15 empty] | 0
0
| 1 | 1 | < 200 | < 100 | < 100 | < 100 | [15 empty] | 1
| 1 | 2 | < 200 | < 100 | < 100 | < 100 | [15 empty] | 10
| 1 | 3 | < 200 | < 100 | < 100 | < 100 | [15 empty] | 16
| 1 | 4 | < 200 | < 100 | < 100 | < 100 | [15 empty] | < 100
| 1 | 5 | < 200 | < 100 | < 100 | < 100 | [15 empty] | 16
| 1 | 6 | < 200 | < 100 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Hg in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 18
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
In | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 2
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 196 (Certified Value 0.038 wt.%) | [20 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.032 | 0.028985 | 0.035156 | 0.04 | 0.04 | 0.05 | 0.03 | 0.039 | 0.05 | 0.04 | 0.0439 | 0.04 | 0.05 | 0.06 | 0.04 | 0.04 | 0.036 |  |  | 1
| 1 | 2 | 0.032 | 0.030015 | 0.035653 | 0.04 | 0.05 | 0.04 | 0.03 | 0.038 | 0.04 | 0.04 | 0.0386 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.036 |  |  | #N/A
| 1 | 3 | 0.031 | 0.028975 | 0.035285 | 0.05 | 0.05 | 0.04 | 0.04 | 0.037 | 0.04 | 0.04 | 0.0422 | 0.04 | 0.05 | 0.06 | 0.04 | 0.04 | 0.036 |  |  | 16
| 1 | 4 | 0.029 | 0.030015 | 0.034785 | 0.04 | 0.04 | 0.04 | 0.03 | 0.041 | 0.04 | 0.04 | 0.0422 | 0.03 | 0.06 | 0.05 | 0.04 | 0.05 | 0.036 |  |  | 0.03834
| 1 | 5 | 0.031 | 0.031055 | 0.036956 | 0.04 | 0.05 | 0.04 | 0.04 | 0.042 | 0.04 | 0.04 | 0.0385 | 0.03 | 0.06 | 0.06 | 0.04 | 0.04 | 0.036 |  |  | 13
| 1 | 6 | 0.034 | 0.0295 | 0.034029 | 0.04 | 0.04 | 0.04 | 0.03 | 0.042 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.06 | 0.04 | 0.04 | 0.036
| Mean |  | 0.0315 | 0.029757 | 0.035311 | 0.041667 | 0.045 | 0.041667 | 0.033333 | 0.039833 | 0.041667 | 0.04 | 0.0409 | 0.036667 | 0.053333 | 0.056667 | 0.04 | 0.041667 | 0.036
| Median |  | 0.0315 | 0.029757 | 0.035221 | 0.04 | 0.045 | 0.04 | 0.03 | 0.04 | 0.04 | 0.04 | 0.0411 | 0.04 | 0.05 | 0.06 | 0.04 | 0.04 | 0.036
| Std Dev. |  | 0.0016432 | 0.00078632 | 0.00097647 | 0.0040825 | 0.0054772 | 0.0040825 | 0.005164 | 0.002137 | 0.0040825 | 0 | 0.0022018 | 0.005164 | 0.005164 | 0.005164 | 0 | 0.0040825 | 0
| Rel.Std.Dev. |  | 5.216% | 2.642% | 2.765% | 9.798% | 12.17% | 9.798% | 15.49% | 5.365% | 9.798% | 0% | 5.383% | 14.08% | 9.682% | 9.113% | 0% | 9.798% | 0%
| PDM3 |  | -17.84% | -22.39% | -7.902% | 8.676% | 17.37% | 8.676% | -13.06% | 3.894% | 8.676% | 4.329% | 6.676% | -4.365% | 39.1% | 47.8% | 4.329% | 8.676% | -6.104%
| Z-Score (Absolute) |  | 1.72 | 2.07 | 0.95 | 0.34 | 1.01 | 0.34 | 1.35 | 0.03 | 0.34 | 0 | 0.18 | 0.67 | 2.7 | 3.37 | 0 | 0.34 | 0.78
| Analytical results for La in OREAS 196 (Indicative Value 68 ppm) | [20 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 07 | 08 | 09 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 90 | < 90 | 90 | [16 empty] | 1
| 1 | 2 | < 90 | < 90 | 90 | [16 empty] | 3
| 1 | 3 | 90 | < 90 | 90 | [16 empty] | 16
| 1 | 4 | < 90 | < 90 | 90 | [16 empty] | 67.5023
| 1 | 5 | 90 | < 90 | 90 | [16 empty] | 9
| 1 | 6 | < 90 | < 90 | < 90
| Mean |  | 90 |  | 90
| Median |  | 90 |  | 90
| Std Dev. |  | 0 |  | 0
| Rel.Std.Dev. |  | 0% |  | 0%
| PDM3 |  | 33.33% |  | 33.33%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.01
| Analytical results for MgO in OREAS 196 (Certified Value 12.67 wt.%) | [20 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 12.67 | 12.74 | 12.7119 | 12.7127 | 12.14 | 12.55 | 12.65 | 12.65 | 12.3579 | 12.77 | 12.08 | 12.7666 | 12.9 | 12.87 | 12.36 | 12.68 | 12.735 | 12.73 | 12.6952 | 1
| 1 | 2 | 12.54 | 12.88 | 12.7666 | 12.7894 | 11.96 | 12.65 | 12.7 | 12.6 | 12.3664 | 12.78 | 12.1 | 12.4647 | 12.9 | 12.85 | 12.35 | 12.71 | 12.752 | 12.78 | 12.6537 | #N/A
| 1 | 3 | 12.47 | 12.81 | 12.7667 | 12.7658 | 12.15 | 12.65 | 12.65 | 12.75 | 12.443 | 12.78 | 12.08 | 12.644 | 12.9 | 12.81 | 12.41 | 12.64 | 12.735 | 12.71 | 12.6703 | 16
| 1 | 4 | 12.34 | 12.71 | 12.7979 | 12.6387 | 12.18 | 12.4 | 12.65 | 12.65 | 12.476 | 12.82 | 12.08 | 12.2523 | 12.8 | 12.84 | 12.36 | 12.68 | 12.735 | 12.75 | 12.6968 | 12.6716
| 1 | 5 | 12.95 | 12.73 | 12.8688 | 12.7959 | 12.04 | 12.65 | 12.65 | 12.6 | 12.3718 | 12.77 | 12.03 | 12.7635 | 12.9 | 12.89 | 12.34 | 12.62 | 12.801 | 12.76 | 12.6919 | 14
| 1 | 6 | 12.74 | 12.75 | 12.7989 | 12.643 | 12.25 | 12.55 | 12.7 | 12.45 | 12.3601 | 12.8 | 12.09 | 12.4731 | 12.8 | 12.83 | 12.39 | 12.61 | 12.718 | 12.73 | 12.725
| Mean |  | 12.6183 | 12.77 | 12.7851 | 12.7243 | 12.12 | 12.575 | 12.6667 | 12.6167 | 12.3959 | 12.7867 | 12.0767 | 12.5607 | 12.8667 | 12.8483 | 12.3683 | 12.6567 | 12.746 | 12.7433 | 12.6888
| Median |  | 12.605 | 12.745 | 12.7823 | 12.7393 | 12.145 | 12.6 | 12.65 | 12.625 | 12.3691 | 12.78 | 12.08 | 12.5586 | 12.9 | 12.845 | 12.36 | 12.66 | 12.735 | 12.74 | 12.6935
| Std Dev. |  | 0.21591 | 0.063561 | 0.051773 | 0.070913 | 0.10373 | 0.098742 | 0.02582 | 0.098319 | 0.050624 | 0.019664 | 0.024221 | 0.20123 | 0.05164 | 0.028577 | 0.026394 | 0.039328 | 0.02901 | 0.025033 | 0.024494
| Rel.Std.Dev. |  | 1.711% | 0.4977% | 0.4049% | 0.5573% | 0.8559% | 0.7852% | 0.2038% | 0.7793% | 0.4084% | 0.1538% | 0.2006% | 1.602% | 0.4013% | 0.2224% | 0.2134% | 0.3107% | 0.2276% | 0.1964% | 0.193%
| PDM3 |  | -0.4205% | 0.7764% | 0.8959% | 0.4154% | -4.353% | -0.7625% | -0.03906% | -0.4336% | -2.176% | 0.9079% | -4.695% | -0.8753% | 1.539% | 1.395% | -2.393% | -0.118% | 0.587% | 0.566% | 0.1358%
| Z-Score (Absolute) |  | 0.32 | 0.67 | 0.77 | 0.38 | 3.57 | 0.6 | 0 | 0.33 | 1.77 | 0.78 | 3.85 | 0.69 | 1.31 | 1.19 | 1.95 | 0.07 | 0.52 | 0.5 | 0.14
| Analytical results for MnO in OREAS 196 (Certified Value 0.494 wt.%) | [20 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 | 26 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.429 | 0.5 | 0.5245 | 0.4966 | 0.484 | 0.504 | 0.491 | 0.478 | 0.429 | 0.5 | 0.484 | 0.4829 | 0.49 | 0.5 | 0.49 | 0.52 | 0.493 | 0.4945 | 0.5 | 1
| 1 | 2 | 0.433 | 0.51 | 0.52408 | 0.49172 | 0.479 | 0.504 | 0.491 | 0.478 | 0.43 | 0.5 | 0.474 | 0.49443 | 0.49 | 0.49 | 0.48 | 0.5 | 0.493 | 0.4984 | 0.498 | #N/A
| 1 | 3 | 0.429 | 0.51 | 0.52502 | 0.49623 | 0.486 | 0.504 | 0.491 | 0.491 | 0.421 | 0.5 | 0.493 | 0.49006 | 0.5 | 0.5 | 0.49 | 0.51 | 0.493 | 0.4945 | 0.498 | 16
| 1 | 4 | 0.424 | 0.51 | 0.52504 | 0.49673 | 0.481 | 0.504 | 0.491 | 0.478 | 0.426 | 0.5 | 0.484 | 0.49034 | 0.49 | 0.53 | 0.48 | 0.52 | 0.502 | 0.4997 | 0.5 | 0.49423
| 1 | 5 | 0.44 | 0.5 | 0.5281 | 0.49219 | 0.48 | 0.504 | 0.491 | 0.478 | 0.417 | 0.5 | 0.484 | 0.49713 | 0.49 | 0.52 | 0.48 | 0.51 | 0.484 | 0.4958 | 0.502 | 15
| 1 | 6 | 0.454 | 0.51 | 0.52697 | 0.49461 | 0.486 | 0.491 | 0.491 | 0.478 | 0.428 | 0.5 | 0.484 | 0.47797 | 0.49 | 0.52 | 0.49 | 0.52 | 0.493 | 0.4945 | 0.501
| Mean |  | 0.43483 | 0.50667 | 0.52562 | 0.49468 | 0.48267 | 0.50183 | 0.491 | 0.48017 | 0.42517 | 0.5 | 0.48383 | 0.4888 | 0.49167 | 0.51 | 0.485 | 0.51333 | 0.493 | 0.49623 | 0.49983
| Median |  | 0.431 | 0.51 | 0.52503 | 0.49542 | 0.4825 | 0.504 | 0.491 | 0.478 | 0.427 | 0.5 | 0.484 | 0.4902 | 0.49 | 0.51 | 0.485 | 0.515 | 0.493 | 0.49515 | 0.5
| Std Dev. |  | 0.010797 | 0.005164 | 0.0015687 | 0.0022497 | 0.0030768 | 0.0053072 | 6.0809e-17 | 0.0053072 | 0.0051153 | 0 | 0.0060139 | 0.0071714 | 0.0040825 | 0.015492 | 0.0054772 | 0.008165 | 0.0056921 | 0.0022765 | 0.0016021
| Rel.Std.Dev. |  | 2.483% | 1.019% | 0.2984% | 0.4548% | 0.6375% | 1.058% | 1.238e-14% | 1.105% | 1.203% | 0% | 1.243% | 1.467% | 0.8303% | 3.038% | 1.129% | 1.591% | 1.155% | 0.4588% | 0.3205%
| PDM3 |  | -12.02% | 2.516% | 6.351% | 0.09075% | -2.34% | 1.538% | -0.6541% | -2.846% | -13.97% | 1.167% | -2.104% | -1.098% | -0.5192% | 3.19% | -1.868% | 3.865% | -0.2494% | 0.4048% | 1.133%
| Z-Score (Absolute) |  | 4.36 | 1.09 | 2.53 | 0.18 | 0.73 | 0.69 | 0.12 | 0.94 | 5.09 | NA | 0.66 | 0.27 | 0.05 | 1.34 | 0.55 | 1.59 | 0.05 | 0.3 | 0.57
Results from laboratory 11 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Mo in OREAS 196 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 07 | 08 | 09 | 15 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 50 | < 50 | 70 | < 100 | [15 empty] | 1
| 1 | 2 | < 50 | < 50 | 60 | < 100 | [15 empty] | 4
| 1 | 3 | < 50 | < 50 | 70 | < 100 | [15 empty] | 16
| 1 | 4 | < 50 | < 50 | 70 | < 100 | [15 empty] | < 50
| 1 | 5 | < 50 | < 50 | 70 | < 100 | [15 empty] | 10
| 1 | 6 | < 50 | < 50 | < 50 | < 100
| Mean |  |  |  | 68
| Median |  |  |  | 70
| Std Dev. |  |  |  | 4.4721
| Rel.Std.Dev. |  |  |  | 6.577%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 1.26 | 0.67
| Analytical results for Na2O in OREAS 196 (Indicative Value 0.115 wt.%) | [20 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.03 | < 0.01 | 0.12714 | 0.06 | 0.068 | 0.13 | 0.13 | 0.1058 | 0.1 | 0.26 | 0.08 | 0.04 | 0.21 | [6 empty] | 1
| 1 | 2 | < 0.01 | < 0.01 | 0.13507 | 0.05 | 0.071 | 0.12 | 0.14 | 0.1104 | 0.14 | 0.26 | 0.09 | 0.04 | 0.21 | [6 empty] | 5
| 1 | 3 | 0.01 | < 0.01 | 0.1181 | 0.05 | 0.069 | 0.14 | 0.14 | 0.1021 | 0.16 | 0.26 | 0.09 | 0.04 | 0.21 | [6 empty] | 16
| 1 | 4 | < 0.01 | < 0.01 | 0.1344 | 0.06 | 0.072 | 0.12 | 0.12 | 0.1299 | 0.14 | 0.25 | 0.09 | 0.04 | 0.21 | [6 empty] | 0.11496
| 1 | 5 | 0.02 | < 0.01 | 0.11739 | 0.04 | 0.071 | 0.13 | 0.13 | 0.1225 | 0.14 | 0.26 | 0.09 | 0.05 | 0.21 | [6 empty] | 11
| 1 | 6 | 0.02 | < 0.01 | 0.13775 | 0.07 | 0.071 | 0.12 | 0.12 | 0.1347 | 0.12 | 0.25 | 0.09 | 0.05 | 0.21
| Mean |  | 0.02 |  | 0.12831 | 0.055 | 0.070333 | 0.12667 | 0.13 | 0.11757 | 0.13333 | 0.25667 | 0.088333 | 0.043333 | 0.21
| Median |  | 0.02 |  | 0.13077 | 0.055 | 0.071 | 0.125 | 0.13 | 0.11645 | 0.14 | 0.26 | 0.09 | 0.04 | 0.21
| Std Dev. |  | 0.008165 |  | 0.0089074 | 0.010488 | 0.0015055 | 0.008165 | 0.0089443 | 0.013409 | 0.020656 | 0.005164 | 0.0040825 | 0.005164 | 0
| Rel.Std.Dev. |  | 40.82% |  | 6.942% | 19.07% | 2.141% | 6.446% | 6.88% | 11.41% | 15.49% | 2.012% | 4.622% | 11.92% | 0%
| PDM3 |  | -82.6% |  | 11.61% | -52.16% | -38.82% | 10.18% | 13.08% | 2.266% | 15.98% | 123.3% | -23.16% | -62.31% | 82.67%
| Z-Score (Absolute) |  | 1.46 | 1.61 | 0.15 | 0.89 | 0.67 | 0.13 | 0.18 | 0 | 0.23 | 1.99 | 0.42 | 1.06 | 1.32
| Analytical results for Nb in OREAS 196 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 07 | 08 | 09 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | < 50 | < 50 | 50 | [16 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | [16 empty] | 6
| 1 | 3 | < 50 | < 50 | < 50 | [16 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | [16 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | [16 empty] | 12
| 1 | 6 | < 50 | < 50 | < 50
| Mean |  |  |  | 50
| Median |  |  |  | 50
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Ni in OREAS 196 (Certified Value 2.09 wt.%) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 18 | 20 | 25 | 26 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 2.016 | 2.13 | 2.1156 | 2.0967 | 2.1037 | 2.04 | 2.09 | 2.06 | 2.132 | 2.01 | 2.0921 | 2.064 | 2.04 | 2.14 | 2.09 | 2.12 | 2.116 |  |  | 1
| 1 | 2 | 2.011 | 2.15 | 2.1213 | 2.0849 | 2.0621 | 2.04 | 2.09 | 2.05 | 2.131 | 2.03 | 2.0616 | 2.056 | 2.04 | 2.12 | 2.08 | 2.13 | 2.101 |  |  | 22
| 1 | 3 | 2.019 | 2.14 | 2.1164 | 2.0989 | 2.1084 | 2.04 | 2.08 | 2.08 | 2.142 | 1.94 | 2.0929 | 2.062 | 2.04 | 2.13 | 2.09 | 2.11 | 2.105 |  |  | 16
| 1 | 4 | 1.97 | 2.14 | 2.1224 | 2.0874 | 2.1029 | 2.03 | 2.09 | 2.06 | 2.142 | 2.04 | 2.1094 | 2.039 | 2.03 | 2.16 | 2.09 | 2.13 | 2.114 |  |  | 2.0858
| 1 | 5 | 2.083 | 2.14 | 2.1368 | 2.092 | 2.0778 | 2.05 | 2.09 | 2.05 | 2.139 | 2.06 | 2.1034 | 2.049 | 2.03 | 2.13 | 2.08 | 2.12 | 2.115 |  |  | 16
| 1 | 6 | 2.151 | 2.15 | 2.13 | 2.0972 | 2.1171 | 2.04 | 2.09 | 2.03 | 2.148 | 2.01 | 2.0848 | 2.058 | 2.03 | 2.13 | 2.09 | 2.11 | 2.114
| Mean |  | 2.0417 | 2.1417 | 2.1237 | 2.0928 | 2.0953 | 2.04 | 2.0883 | 2.055 | 2.139 | 2.015 | 2.0907 | 2.0547 | 2.035 | 2.135 | 2.0867 | 2.12 | 2.1108
| Median |  | 2.0175 | 2.14 | 2.1219 | 2.0943 | 2.1033 | 2.04 | 2.09 | 2.055 | 2.1405 | 2.02 | 2.0925 | 2.057 | 2.035 | 2.13 | 2.09 | 2.12 | 2.114
| Std Dev. |  | 0.06467 | 0.0075277 | 0.0081937 | 0.0057405 | 0.020898 | 0.0063246 | 0.0040825 | 0.016432 | 0.0065115 | 0.041352 | 0.01673 | 0.009288 | 0.0054772 | 0.013784 | 0.005164 | 0.0089443 | 0.0062423
| Rel.Std.Dev. |  | 3.168% | 0.3515% | 0.3858% | 0.2743% | 0.9974% | 0.31% | 0.1955% | 0.7996% | 0.3044% | 2.052% | 0.8002% | 0.452% | 0.2692% | 0.6456% | 0.2475% | 0.4219% | 0.2957%
| PDM3 |  | -2.116% | 2.678% | 1.819% | 0.3379% | 0.457% | -2.196% | 0.1214% | -1.477% | 2.551% | -3.394% | 0.2349% | -1.493% | -2.436% | 2.359% | 0.04153% | 1.64% | 1.2%
| Z-Score (Absolute) |  | 0.93 | 0.96 | 0.62 | 0.04 | 0.09 | 0.96 | 0.04 | 0.67 | 0.91 | 1.43 | 0 | 0.68 | 1.05 | 0.84 | 0.08 | 0.55 | 0.38
| Analytical results for P2O5 in OREAS 196 (Certified Value 0.008 wt.%) | [20 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | < 0.005 | 0.007 | < 0.004 | 0.0042614 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.007 | < 0.01 | 0.00802 | 0.0092 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.0092 |  | 1
| 1 | 2 | < 0.005 | 0.006 | < 0.004 | 0.0071306 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.006 | < 0.01 | 0.00458 | 0.0046 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0115 |  | #N/A
| 1 | 3 | < 0.005 | 0.006 | < 0.004 | 0.0074847 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.007 | < 0.01 | 0.00756 | 0.0046 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.0115 |  | 16
| 1 | 4 | < 0.005 | 0.007 | < 0.004 | 0.0063246 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.008 | < 0.01 | 0.00229 | 0.0137 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.0092 |  | 0.0083227
| 1 | 5 | < 0.005 | 0.006 | < 0.004 | 0.0042236 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.005 | < 0.01 | 0.00435 | 0.0115 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0115 |  | 17
| 1 | 6 | < 0.005 | 0.007 | < 0.004 | 0.005317 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.008 | < 0.01 | 0.00871 | 0.0115 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0115
| Mean |  |  | 0.0065 |  | 0.0057903 | 0.01 |  | 0.01 |  |  | 0.0068333 |  | 0.0059183 | 0.0091833 |  |  | 0.01 |  | 0.010733
| Median |  |  | 0.0065 |  | 0.0058208 | 0.01 |  | 0.01 |  |  | 0.007 |  | 0.00607 | 0.01035 |  |  | 0.01 |  | 0.0115
| Std Dev. |  |  | 0.00054772 |  | 0.001412 | 0 |  | 0 |  |  | 0.001169 |  | 0.0025425 | 0.0038249 |  |  | 0 |  | 0.0011877
| Rel.Std.Dev. |  |  | 8.427% |  | 24.39% | 0% |  | 0% |  |  | 17.11% |  | 42.96% | 41.65% |  |  | 0% |  | 11.07%
| PDM3 |  |  | -21.9% |  | -30.43% | 20.15% |  | 20.15% |  |  | -17.9% |  | -28.89% | 10.34% |  |  | 20.15% |  | 28.96%
| Z-Score (Absolute) |  | 2.4 | 0.91 | 2.81 | 0.33 | 3.12 | 0.33 | 3.81 | 0.33 | 0.33 | 1.19 | 0.33 | 0.43 | 3.12 | 0.33 | 0.33 | 1.74 | 0.33 | 4.39
| Analytical results for Pb in OREAS 196 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 13 | 14 | 15 | 20 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 190 | < 50 | < 50 | < 50 | < 50 | 154 | < 10 | < 100 | < 90 | [10 empty] | 1
| 1 | 2 | < 190 | < 50 | < 50 | < 50 | < 50 | 151 | 10 | < 100 | < 90 | [10 empty] | 7
| 1 | 3 | < 190 | < 50 | < 50 | < 50 | < 50 | 155 | 10 | < 100 | < 90 | [10 empty] | 16
| 1 | 4 | < 190 | < 50 | 50 | < 50 | < 50 | 308 | 20 | < 100 | < 90 | [10 empty] | < 50
| 1 | 5 | < 190 | < 50 | < 50 | < 50 | < 50 | 169 | 30 | < 100 | < 90 | [10 empty] | 13
| 1 | 6 | < 190 | < 50 | < 50 | < 50 | < 50 | 156 | 10 | < 100 | < 90
| Mean |  |  |  | 50 |  |  | 182.1667 | 16
| Median |  |  |  | 50 |  |  | 155.5 | 10
| Std Dev. | [6 empty] | 61.9594 | 8.9443
| Rel.Std.Dev. | [6 empty] | 34.01% | 55.9%
PDM3
| Z-Score (Absolute) |  | 2.86 | 0.19 | 0 | 0.19 | 0.19 | 6.88 | 0.67 | 0.94 | 0.78
| Analytical results for Rb in OREAS 196 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 07 | 08 | 09 | 15 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 8
| 1 | 3 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 100 | [15 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 14
| 1 | 6 | < 50 | < 50 | < 50 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Sb in OREAS 196 (Indicative Value < 50 ppm) | [20 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 07 | 08 | 09 | 15 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 9
| 1 | 3 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 100 | [15 empty] | < 50
| 1 | 5 | < 50 | < 50 | < 50 | < 100 | [15 empty] | 15
| 1 | 6 | < 50 | < 50 | < 50 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Sc in OREAS 196 (Indicative Value 20.8 ppm) | [20 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 15 | 18 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | [17 empty] | 1
1
| 1 | 1 | < 100 | 20 | [17 empty] | 1
| 1 | 2 | < 100 | 21 | [17 empty] | 10
| 1 | 3 | < 100 | 21 | [17 empty] | 16
| 1 | 4 | < 100 | 24 | [17 empty] | 20.8333
| 1 | 5 | < 100 | 23 | [17 empty] | 16
| 1 | 6 | < 100 | 16
| Mean |  |  | 20.8333
| Median |  |  | 21
| Std Dev. |  |  | 2.7869
| Rel.Std.Dev. |  |  | 13.38%
| PDM3 |  |  | 1.554e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Se in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 1
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 196 (Certified Value 41.98 wt.%) | [20 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 18 | 20 | 25 | 26 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 40.39 | 42.74 | 40.8003 | 41.9607 | 40.44 | 42 | 42.4 | 42.5 | 39.631 | 42.63 | 40.74 | 42.54 | 39.36 | 40.27 | 42.1 | 42.35 | 42.28 | 42.284 |  | 1
| 1 | 2 | 40 | 42.33 | 40.9183 | 42.0851 | 40.01 | 42 | 42.4 | 42.3 | 39.536 | 42.52 | 40.65 | 42.43 | 39.36 | 40.43 | 42.1 | 42.37 | 42.32 | 42.267 |  | #N/A
| 1 | 3 | 40.26 | 42.53 | 40.9059 | 42.0013 | 40.77 | 42.2 | 42.2 | 42.8 | 39.628 | 42.67 | 40.52 | 42.43 | 39.39 | 40.36 | 42 | 42.26 | 42.21 | 42.207 |  | 16
| 1 | 4 | 39.63 | 42.65 | 40.9654 | 42.0601 | 40.38 | 41.3 | 42.2 | 42.3 | 39.722 | 42.57 | 40.66 | 42.35 | 39.38 | 40.46 | 42.5 | 42.36 | 42.36 | 42.357 |  | 41.9771
| 1 | 5 | 41.32 | 42.83 | 40.8611 | 41.9292 | 40.1 | 42.1 | 42.4 | 42.4 | 39.622 | 42.66 | 40.31 | 42.42 | 39.39 | 40.4 | 42 | 42.36 | 42.38 | 42.266 |  | 18
| 1 | 6 | 42.33 | 42.64 | 41.007 | 41.9934 | 40.69 | 42.2 | 42.3 | 41.8 | 39.635 | 42.73 | 40.71 | 42.16 | 39.4 | 40.38 | 42 | 42.2 | 42.33 | 42.344
| Mean |  | 40.655 | 42.62 | 40.9097 | 42.005 | 40.3983 | 41.9667 | 42.3167 | 42.35 | 39.629 | 42.63 | 40.5983 | 42.3883 | 39.38 | 40.3833 | 42.1167 | 42.3167 | 42.3133 | 42.2875
| Median |  | 40.325 | 42.645 | 40.9121 | 41.9973 | 40.41 | 42.05 | 42.35 | 42.35 | 39.6295 | 42.645 | 40.655 | 42.425 | 39.385 | 40.39 | 42.05 | 42.355 | 42.325 | 42.2755
| Std Dev. |  | 0.99536 | 0.17436 | 0.073491 | 0.058866 | 0.30499 | 0.33862 | 0.098319 | 0.32711 | 0.058971 | 0.0751 | 0.16018 | 0.12734 | 0.016733 | 0.065929 | 0.19408 | 0.070048 | 0.06121 | 0.055493
| Rel.Std.Dev. |  | 2.448% | 0.4091% | 0.1796% | 0.1401% | 0.7549% | 0.8069% | 0.2323% | 0.7724% | 0.1488% | 0.1762% | 0.3945% | 0.3004% | 0.04249% | 0.1633% | 0.4608% | 0.1655% | 0.1447% | 0.1312%
| PDM3 |  | -3.149% | 1.532% | -2.543% | 0.0665% | -3.761% | -0.02476% | 0.809% | 0.8884% | -5.594% | 1.555% | -3.284% | 0.9798% | -6.187% | -3.797% | 0.3326% | 0.809% | 0.8011% | 0.7396%
| Z-Score (Absolute) |  | 2.14 | 0.85 | 1.75 | 0.08 | 2.53 | 0.14 | 0.39 | 0.44 | 3.7 | 0.87 | 2.22 | 0.5 | 4.08 | 2.55 | 0.08 | 0.39 | 0.38 | 0.34
| Analytical results for Sn in OREAS 196 (Certified Value < 50 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 13 | 14 | 15 | 20 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 160 | < 50 | < 50 | < 50 | < 50 | 64 | < 10 | < 100 | < 80 | [10 empty] | 1
| 1 | 2 | < 160 | < 50 | < 50 | < 50 | 50 | 15 | < 10 | < 100 | < 80 | [10 empty] | 15
| 1 | 3 | < 160 | < 50 | < 50 | < 50 | 50 | 89 | 160 | < 100 | < 80 | [10 empty] | 16
| 1 | 4 | < 160 | < 50 | < 50 | 60 | 50 |  | < 10 | < 100 | < 80 | [10 empty] | < 50
| 1 | 5 | < 160 | < 50 | < 50 | < 50 | 60 | 62 | 70 | < 100 | < 80 | [10 empty] | 19
| 1 | 6 | < 160 | < 50 | < 50 | < 50 | < 50 | 38 | 60 | < 100 | < 80
| Mean |  |  |  |  | 60 | 52.5 | 53.6 | 96.6667
| Median |  |  |  |  | 60 | 50 | 62 | 70
| Std Dev. | [5 empty] | 5 | 28.1301 | 55.0757
| Rel.Std.Dev. | [5 empty] | 9.524% | 52.48% | 56.97%
PDM3
| Z-Score (Absolute) |  | 2.33 | 1.2 | 1.2 | 0.82 | 0 | 0.67 | 0.49 | 0.44 | 0.26
| Analytical results for SO3 in OREAS 196 (Certified Value 0.504 wt.%) | [20 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 18 | 20 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.441 | 0.5744 | 0.51734 | 0.43679 | 0.499 | 0.499 | 0.524 | 0.506 | 0.524 | 0.45199 | 0.482 | 0.54 | 0.09 | 0.51 | [5 empty] | 1
| 1 | 2 | 0.438 | 0.5818 | 0.52011 | 0.42376 | 0.499 | 0.499 | 0.524 | 0.508 | 0.521 | 0.49295 | 0.4919 | 0.55 | 0.08 | 0.51 | [5 empty] | #N/A
| 1 | 3 | 0.449 | 0.5793 | 0.51898 | 0.41166 | 0.499 | 0.499 | 0.549 | 0.521 | 0.524 | 0.45649 | 0.4845 | 0.55 | 0.09 | 0.5 | [5 empty] | 16
| 1 | 4 | 0.433 | 0.6043 | 0.52454 | 0.41953 | 0.499 | 0.524 | 0.549 | 0.516 | 0.523 | 0.50693 | 0.4745 | 0.54 | 0.09 | 0.51 | [5 empty] | 0.50394
| 1 | 5 | 0.462 | 0.5744 | 0.52228 | 0.4308 | 0.499 | 0.499 | 0.549 | 0.509 | 0.519 | 0.41304 | 0.487 | 0.54 | 0.1 | 0.51 | [5 empty] | 20
| 1 | 6 | 0.477 | 0.5818 | 0.5212 | 0.4126 | 0.499 | 0.499 | 0.524 | 0.511 | 0.525 | 0.4892 | 0.4745 | 0.54 | 0.09 | 0.5
| Mean |  | 0.45 | 0.58267 | 0.52074 | 0.42252 | 0.499 | 0.50317 | 0.5365 | 0.51183 | 0.52267 | 0.46843 | 0.4824 | 0.54333 | 0.09 | 0.50667
| Median |  | 0.445 | 0.58055 | 0.52065 | 0.42165 | 0.499 | 0.499 | 0.5365 | 0.51 | 0.5235 | 0.47285 | 0.48325 | 0.54 | 0.09 | 0.51
| Std Dev. |  | 0.016661 | 0.011113 | 0.0025343 | 0.009991 | 6.0809e-17 | 0.010206 | 0.013693 | 0.0056362 | 0.0022509 | 0.034629 | 0.0069397 | 0.005164 | 0.0063246 | 0.005164
| Rel.Std.Dev. |  | 3.703% | 1.907% | 0.4867% | 2.365% | 1.219e-14% | 2.028% | 2.552% | 1.101% | 0.4307% | 7.392% | 1.439% | 0.9504% | 7.027% | 1.019%
| PDM3 |  | -10.7% | 15.62% | 3.334% | -16.16% | -0.9797% | -0.1529% | 6.462% | 1.567% | 3.717% | -7.045% | -4.274% | 7.818% | -82.14% | 0.5417%
| Z-Score (Absolute) |  | 1.36 | 1.92 | 0.39 | 2.04 | 0.14 | 0.04 | 0.79 | 0.17 | 0.43 | 0.91 | 0.56 | 0.94 | 10.27 | 0.04
| Analytical results for SrO in OREAS 196 (Indicative Value < 120 ppm) | [20 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 07 | 08 | 09 | 13 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 59 | < 120 | < 120 | < 120 | 34.3 | 24 | < 100 | < 100 | [9 empty] | 1
| 1 | 2 | < 50 | < 50 | < 59 | < 120 | < 120 | < 120 | 61.5 | 35 | < 100 | < 100 | [9 empty] | 2
| 1 | 3 | < 50 | < 50 | < 59 | < 120 | < 120 | < 120 | 54.4 | 35 | < 100 | < 100 | [9 empty] | 16
| 1 | 4 | < 50 | < 50 | < 59 | < 120 | < 120 | < 120 |  | 35 | < 100 | < 100 | [9 empty] | < 120
| 1 | 5 | < 50 | < 50 | < 59 | < 120 | < 120 | < 120 | 133.6 | 35 | < 100 | < 100 | [9 empty] | 8
| 1 | 6 | < 50 | < 50 | < 59 | < 120 | < 120 | < 120 | 156.1 | 35 | < 100 | < 100
| Mean | [7 empty] | 87.98 | 33.1667
| Median | [7 empty] | 61.5 | 35
| Std Dev. | [7 empty] | 53.46 | 4.4907
| Rel.Std.Dev. | [7 empty] | 60.76% | 13.54%
PDM3
| Z-Score (Absolute) |  | 1.32 | 1.32 | 1.08 | 0.48 | 0.48 | 0.48 | 2 | 0.87 | 0 | 0
| Analytical results for Ta in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ta | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 3
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 9
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 4
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 10
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 196 (Certified Value 0.139 wt.%) | [20 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 18 | 20 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.1 | 0.13 | 0.12069 | 0.1321 | 0.14 | 0.13 | 0.14 | 0.17 | 0.151 | 0.14 | 0.14 | 0.14 | 0.13 | 0.13 | 0.16 | 0.14 |  |  |  | 1
| 1 | 2 | 0.1 | 0.14 | 0.12071 | 0.12835 | 0.14 | 0.13 | 0.14 | 0.16 | 0.148 | 0.14 | 0.14 | 0.14 | 0.13 | 0.13 | 0.15 | 0.15 |  |  |  | #N/A
| 1 | 3 | 0.11 | 0.13 | 0.12169 | 0.1429 | 0.14 | 0.13 | 0.14 | 0.17 | 0.144 | 0.15 | 0.14 | 0.14 | 0.12 | 0.13 | 0.15 | 0.14 |  |  |  | 16
| 1 | 4 | 0.1 | 0.14 | 0.12487 | 0.13282 | 0.14 | 0.13 | 0.14 | 0.18 | 0.148 | 0.14 | 0.14 | 0.14 | 0.12 | 0.13 | 0.15 | 0.15 |  |  |  | 0.1388
| 1 | 5 | 0.11 | 0.13 | 0.11963 | 0.13093 | 0.14 | 0.13 | 0.14 | 0.18 | 0.153 | 0.14 | 0.14 | 0.15 | 0.11 | 0.13 | 0.15 | 0.14 |  |  |  | 21
| 1 | 6 | 0.11 | 0.13 | 0.11916 | 0.13931 | 0.14 | 0.13 | 0.14 | 0.15 | 0.15 | 0.14 | 0.14 | 0.14 | 0.13 | 0.13 | 0.15 | 0.14
| Mean |  | 0.105 | 0.13333 | 0.12112 | 0.1344 | 0.14 | 0.13 | 0.14 | 0.16833 | 0.149 | 0.14167 | 0.14 | 0.14167 | 0.12333 | 0.13 | 0.15167 | 0.14333
| Median |  | 0.105 | 0.13 | 0.1207 | 0.13246 | 0.14 | 0.13 | 0.14 | 0.17 | 0.149 | 0.14 | 0.14 | 0.14 | 0.125 | 0.13 | 0.15 | 0.14
| Std Dev. |  | 0.0054772 | 0.005164 | 0.0020399 | 0.005526 | 0 | 0 | 0 | 0.01169 | 0.0030984 | 0.0040825 | 0 | 0.0040825 | 0.008165 | 0 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 5.216% | 3.873% | 1.684% | 4.112% | 0% | 0% | 0% | 6.945% | 2.079% | 2.882% | 0% | 2.882% | 6.62% | 0% | 2.692% | 3.603%
| PDM3 |  | -24.35% | -3.94% | -12.74% | -3.171% | 0.8634% | -6.341% | 0.8634% | 21.28% | 7.347% | 2.064% | 0.8634% | 2.064% | -11.14% | -6.341% | 9.269% | 3.265%
| Z-Score (Absolute) |  | 3.01 | 0.57 | 1.62 | 0.48 | 0 | 0.86 | 0 | 2.44 | 0.77 | 0.14 | 0 | 0.14 | 1.43 | 0.86 | 1 | 0.29
| Analytical results for Tl in OREAS 196 (Indicative Value < 100 ppm) | [20 empty] | Indicative
Tl | Round | Replicate | Lab | [18 empty] | 1
| No | No | 15 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 5
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 11
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 196 (Certified Value 171 ppm) | [20 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 20 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 140 | < 200 | 181.1087 | 170 | 180 | < 180 | 180 | < 100 | 200 | 107.1 | 179 | < 180 | 200 | [6 empty] | 1
| 1 | 2 | 160 | < 200 | 162.9843 | 170 | 180 | < 180 | 180 | < 100 | 200 | 133.9 | 143 | < 180 | 200 | [6 empty] | #N/A
| 1 | 3 | 120 | < 200 | 160.386 | 180 | 180 | 180 | 180 | < 100 | 200 | 153.5 | 161 | < 180 | 200 | [6 empty] | 16
| 1 | 4 | 120 | < 200 | 179.1968 | 180 | < 180 | < 180 | 180 | < 100 | 200 | 178.5 | 179 | < 180 | 200 | [6 empty] | 171.3033
| 1 | 5 | 170 | < 200 | 148.8257 | 180 | 180 | < 180 | 180 | < 100 | 100 | 180.3 | 161 | < 180 | 200 | [6 empty] | 22
| 1 | 6 | 150 | < 200 | 170.1453 | 180 | 180 | < 180 | 180 | < 100 | 200 | 164.2 | 125 | < 180 | 200
| Mean |  | 143.3333 |  | 167.1078 | 176.6667 | 180 | 180 | 180 |  | 183.3333 | 152.9167 | 158 |  | 200
| Median |  | 145 |  | 166.5648 | 180 | 180 | 180 | 180 |  | 200 | 158.85 | 161 |  | 200
| Std Dev. |  | 20.6559 |  | 12.2296 | 5.164 | 0 |  | 0 |  | 40.8248 | 28.2588 | 21.0428 |  | 0
| Rel.Std.Dev. |  | 14.41% |  | 7.318% | 2.923% | 0% |  | 0% |  | 22.27% | 18.48% | 13.32% |  | 0%
| PDM3 |  | -16.33% |  | -2.449% | 3.131% | 5.077% | 5.077% | 5.077% |  | 7.023% | -10.73% | -7.766% |  | 16.75%
| Z-Score (Absolute) |  | 0.27 | 1.5 | 0.4 | 0.67 | 0.3 | 1.39 | 0.73 | 2.92 | NA | 0 | 0.14 | 1.81 | NA
Results from laboratories 12 and 20 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 196 (Indicative Value < 10 ppm) | [20 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 07 | 08 | 09 | 15 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 1
1
| 1 | 1 | 10 | < 10 | 10 | < 100 | [15 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | < 100 | [15 empty] | 6
| 1 | 3 | < 10 | < 10 | 20 | < 100 | [15 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | < 100 | [15 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 10 | < 100 | [15 empty] | 12
| 1 | 6 | < 10 | < 10 | < 10 | < 100
| Mean |  | 10 |  | 15
| Median |  | 10 |  | 15
| Std Dev. |  |  |  | 7.0711
| Rel.Std.Dev. |  |  |  | 47.14%
PDM3
| Z-Score (Absolute) |  | 0.51 | 0.84 | 0.51 | 17.36
| Analytical results for Y in OREAS 196 (Indicative Value < 39 ppm) | [20 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 07 | 08 | 09 | [16 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 1
1
| 1 | 1 | < 39 | < 39 | 39 | [16 empty] | 1
| 1 | 2 | < 39 | < 39 | < 39 | [16 empty] | 7
| 1 | 3 | < 39 | < 39 | 39 | [16 empty] | 16
| 1 | 4 | 39 | < 39 | < 39 | [16 empty] | < 39
| 1 | 5 | < 39 | < 39 | < 39 | [16 empty] | 13
| 1 | 6 | < 39 | < 39 | < 39
| Mean |  | 39 |  | 39
| Median |  | 39 |  | 39
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 0.67
| Analytical results for Zn in OREAS 196 (Certified Value 432 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 18 | 20 | 26 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 439.026 | 394.1778 | 450 | 440 | 430 | 430 | 434 | 480 | 300 | 500 | 400 | 416.9345 | [7 empty] | 1
| 1 | 2 | 450.273 | 417.6474 | 440 | 430 | 430 | 430 | 434 | 470 | 300 | 400 | 400 | 413.7211 | [7 empty] | #N/A
| 1 | 3 | 446.658 | 384.2058 | 440 | 430 | 440 | 440 | 420 | 480 | 300 | 400 | 400 | 413.7211 | [7 empty] | 16
| 1 | 4 | 442.641 | 405.9126 | 450 | 430 | 440 | 440 | 442 | 460 | 300 | 400 | 400 | 418.5411 | [7 empty] | 431.5263
| 1 | 5 | 454.692 | 389.1645 | 450 | 430 | 430 | 450 | 445 | 480 | 300 | 400 | 400 | 416.9345 | [7 empty] | 23
| 1 | 6 | 450.273 | 404.0951 | 430 | 440 | 420 | 430 | 431 | 460 | 300 | 400 | 400 | 416.9345
| Mean |  | 447.2605 | 399.2005 | 443.3333 | 433.3333 | 431.6667 | 436.6667 | 434.3333 | 471.6667 | 300 | 416.6667 | 400 | 416.1311
| Median |  | 448.4655 | 399.1364 | 445 | 430 | 430 | 435 | 434 | 475 | 300 | 400 | 400 | 416.9345
| Std Dev. |  | 5.7077 | 12.3295 | 8.165 | 5.164 | 7.5277 | 8.165 | 8.8242 | 9.8319 | 0 | 40.8248 | 0 | 1.9678
| Rel.Std.Dev. |  | 1.276% | 3.089% | 1.842% | 1.192% | 1.744% | 1.87% | 2.032% | 2.085% | 0% | 9.798% | 0% | 0.4729%
| PDM3 |  | 3.646% | -7.491% | 2.736% | 0.4187% | 0.03252% | 1.191% | 0.6505% | 9.302% | -30.48% | -3.444% | -7.306% | -3.568%
| Z-Score (Absolute) |  | 0.79 | 2.04 | 0.56 | 0.03 | 0.13 | 0.17 | 0.03 | 2.23 | NA | NA | 1.89 | 1.04
Results from laboratories 15 and 18 were removed due to their 100 ppm reading resolution.
| Analytical results for Zr in OREAS 196 (Indicative Value 58 ppm) | [20 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 13 | 14 | 15 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 370 | < 50 | < 100 | 100 | 100 |  | 50 | < 100 | [11 empty] | 1
| 1 | 2 | < 370 | < 50 | < 100 | < 100 | 100 | 20 | 50 | < 100 | [11 empty] | 8
| 1 | 3 | < 370 | < 50 | < 100 | < 100 | 100 | 43 | 50 | < 100 | [11 empty] | 16
| 1 | 4 | < 370 | < 50 | < 100 | < 100 | 100 |  | 50 | < 100 | [11 empty] | 57.5833
| 1 | 5 | < 370 | < 50 | < 100 | 100 | < 100 | 3 | 50 | < 100 | [11 empty] | 14
| 1 | 6 | < 370 | < 50 | < 100 | < 100 | 100 |  | 50 | < 100
| Mean |  |  |  |  | 100 | 100 | 22 | 50
| Median |  |  |  |  | 100 | 100 | 20 | 50
| Std Dev. |  |  |  |  | 0 | 0 | 20.0749 | 0
| Rel.Std.Dev. |  |  |  |  | 0% | 0% | 91.25% | 0%
| PDM3 |  |  |  |  | 73.66% | 73.66% | -61.79% | -13.17%
| Z-Score (Absolute) |  | 4.37 | 0.81 | 0 | 0.54 | 1.35 | 0.91 | 0 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C704: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C722: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 196 (Indicative Value 5.09 wt.%) | [20 empty] | Indicative
H2O- | Round | Replicate | Lab | [18 empty] | 1
| No | No | 11 | [18 empty] | wt.%
|  |  | LOI*TGA | [18 empty] | 2
2
| 1 | 1 | 4.99 | [18 empty] | 1
| 1 | 2 | 5.31 | [18 empty] | 12
| 1 | 3 | 5.18 | [18 empty] | 16
| 1 | 4 | 4.94 | [18 empty] | 5.0867
| 1 | 5 | 5.23 | [18 empty] | 18
| 1 | 6 | 4.87
| Mean |  | 5.0867
| Median |  | 5.085
| Std Dev. |  | 0.17716
| Rel.Std.Dev. |  | 3.483%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 196 (Certified Value 9.51 wt.%) | [20 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 20 | 25 | 26 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | LOI*Furnace | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | BF*XRF | BF*XRF | LOI*TGA | LOI*TGA | LOI*Furnace | 2
2
| 1 | 1 | 11.58 | 7.7245 | 8.2372 | 8.797 | 11.8 | 8.25 | 7.68 | 8.94 | 12.029 | 7.67 | 12.0352 | 11.17 | 7.96 | 11.98 | 12.09 | 8.71 | 7.9 | 7.63 | 8.518 | 1
| 1 | 2 | 11.69 | 7.5996 | 8.2889 | 8.987 | 12.02 | 8.31 | 7.68 | 9.05 | 12.0506 | 7.77 | 11.9612 | 11.02 | 7.82 | 11.92 | 12.05 | 8.87 | 7.91 | 7.6 | 8.447 | #N/A
| 1 | 3 | 11.65 | 7.6369 | 8.3225 | 8.7585 | 11.9 | 8.34 | 7.64 | 9.12 | 11.8081 | 7.71 | 11.9517 | 10.39 | 7.82 | 11.95 | 12.11 | 8.72 | 7.89 | 7.61 | 8.436 | 16
| 1 | 4 | 11.82 | 7.6005 | 8.234 | 8.835 | 11.94 | 8.36 | 7.65 | 8.95 | 11.7917 | 7.7 | 11.8564 | 10.57 | 7.87 | 11.91 | 12.14 | 8.74 | 7.91 | 7.63 | 8.542 | 9.5076
| 1 | 5 | 11.67 | 7.6159 | 8.2039 | 8.949 | 11.99 | 8.37 | 7.72 | 8.97 | 11.99 | 7.77 | 12.9761 | 10.33 | 7.95 | 11.99 | 12.07 | 8.8 | 7.9 | 7.56 | 8.489 | 25
| 1 | 6 | 11.74 | 7.5941 | 8.36 | 8.952 | 11.91 | 8.39 | 7.63 | 8.95 | 11.9069 | 7.7 | 12.0812 | 11.04 | 7.82 | 11.93 | 12.16 | 8.84 | 7.9 | 7.55 | 8.535
| Mean |  | 11.6917 | 7.6286 | 8.2744 | 8.8798 | 11.9267 | 8.3367 | 7.6667 | 8.9967 | 11.9294 | 7.72 | 12.1436 | 10.7533 | 7.8733 | 11.9467 | 12.1033 | 8.78 | 7.9017 | 7.5967 | 8.4945
| Median |  | 11.68 | 7.6082 | 8.263 | 8.892 | 11.925 | 8.35 | 7.665 | 8.96 | 11.9484 | 7.705 | 11.9982 | 10.795 | 7.845 | 11.94 | 12.1 | 8.77 | 7.9 | 7.605 | 8.5035
| Std Dev. |  | 0.081833 | 0.049501 | 0.059723 | 0.09494 | 0.077374 | 0.050465 | 0.033267 | 0.072572 | 0.11178 | 0.040988 | 0.41504 | 0.36653 | 0.066232 | 0.03266 | 0.041793 | 0.066633 | 0.0075277 | 0.034448 | 0.045072
| Rel.Std.Dev. |  | 0.6999% | 0.6489% | 0.7218% | 1.069% | 0.6487% | 0.6053% | 0.4339% | 0.8067% | 0.937% | 0.5309% | 3.418% | 3.409% | 0.8412% | 0.2734% | 0.3453% | 0.7589% | 0.09527% | 0.4535% | 0.5306%
| PDM3 |  | 22.97% | -19.76% | -12.97% | -6.603% | 25.44% | -12.32% | -19.36% | -5.374% | 25.47% | -18.8% | 27.73% | 13.1% | -17.19% | 25.65% | 27.3% | -7.652% | -16.89% | -20.1% | -10.66%
| Z-Score (Absolute) |  | 1.76 | 0.7 | 0.31 | 0.06 | 1.91 | 0.27 | 0.67 | 0.13 | 1.91 | 0.64 | 2.04 | 1.2 | 0.55 | 1.92 | 2.01 | 0 | 0.53 | 0.72 | 0.17
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 196 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 12 | 15 | 17 | 20 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES/MS | PF*MS | [11 empty] | 3
3
| 1 | 1 | < 1 | < 0.5 | 0.9 | < 10 | [11 empty] | 1
| 1 | 2 | < 1 | < 0.5 | 2.35 | < 10 | [11 empty] | 14
| 1 | 3 | < 1 | < 0.5 | 1.4333 | 10 | [11 empty] | 16
| 1 | 4 | < 1 | < 0.5 | 2.275 | < 10 | [11 empty] | < 0.5
| 1 | 5 | < 1 | < 0.5 | 2.82 | < 10 | [11 empty] | 20
| 1 | 6 | < 1 | < 0.5 | 1.25 | < 10
| Mean |  |  |  | 1.8381 | 10
| Median |  |  |  | 1.8542 | 10
| Std Dev. |  |  |  | 0.74924
| Rel.Std.Dev. |  |  |  | 40.76%
PDM3
| Z-Score (Absolute) |  | 0.57 | 0.78 | 0.57 | 3.96
| Analytical results for Al in OREAS 196 (Certified Value 1.83 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES |  | 2
3
| 1 | 1 | 1.84 | 1.808 | 1.9066 | 1.83 | 1.831 | 1.821 | 1.821 | 2.1923 | 1.85 | 1.85 | 1.82 | 1.66 | 1.994 | 1.81 |  | 1
| 1 | 2 | 1.86 | 1.8151 | 1.7584 | 1.82 | 1.868 | 1.762 | 1.81 | 2.1873 | 1.84 | 1.85 | 1.83 | 1.55 | 1.886 | 2.01 |  | #N/A
| 1 | 3 | 1.85 | 1.7815 | 1.8856 | 1.82 | 1.842 | 1.789 | 1.842 | 2.1904 | 1.79 | 1.81 | 1.84 | 1.68 | 1.744 | 1.9 |  | 16
| 1 | 4 | 1.88 | 1.8079 | 1.9256 | 1.85 | 1.879 | 1.831 | 1.858 | 2.1914 | 1.82 | 1.86 | 1.81 | 1.67 | 1.88 | 1.9 |  | 1.8317
| 1 | 5 | 1.88 | 1.8053 | 1.9012 | 1.81 | 1.821 | 1.799 | 1.852 | 2.1856 | 1.82 | 1.86 | 1.8 | 1.74 | 1.794 | 1.81 |  | 27
| 1 | 6 | 1.86 | 1.8007 | 1.7458 | 1.82 | 1.821 | 1.821 | 1.831 | 2.1882 | 1.83 | 1.85 | 1.79 | 1.66 | 1.754 | 1.86
| Mean |  | 1.8617 | 1.8031 | 1.8539 | 1.825 | 1.8437 | 1.8038 | 1.8357 | 2.1892 | 1.825 | 1.8467 | 1.815 | 1.66 | 1.842 | 1.8817
| Median |  | 1.86 | 1.8066 | 1.8934 | 1.82 | 1.8365 | 1.81 | 1.8365 | 2.1893 | 1.825 | 1.85 | 1.815 | 1.665 | 1.837 | 1.88
| Std Dev. |  | 0.016021 | 0.01158 | 0.079963 | 0.013784 | 0.024622 | 0.025771 | 0.018446 | 0.0026045 | 0.020736 | 0.018619 | 0.018708 | 0.061644 | 0.096058 | 0.074677
| Rel.Std.Dev. |  | 0.8606% | 0.6423% | 4.313% | 0.7553% | 1.336% | 1.429% | 1.005% | 0.119% | 1.136% | 1.008% | 1.031% | 3.714% | 5.215% | 3.969%
| PDM3 |  | 1.634% | -1.564% | 1.209% | -0.3677% | 0.6513% | -1.523% | 0.2146% | 19.51% | -0.3677% | 0.8151% | -0.9137% | -9.376% | 0.5603% | 2.726%
| Z-Score (Absolute) |  | 0.81 | 1.27 | 0.54 | 0.49 | 0.17 | 1.24 | 0.11 | 12.44 | 0.49 | 0.28 | 0.84 | 6.35 | 0.11 | 1.52
| Analytical results for As in OREAS 196 (Indicative Value < 100 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 12 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS |  |  |  |  | 0
0
| 1 | 1 | < 20 | < 15 | < 20 | < 5 | 100 | < 100 | < 100 | < 5 | 12.6 | 1.25 | < 10 |  |  |  |  | 1
| 1 | 2 | < 20 | < 15 | < 20 | 5 | 100 | < 100 | < 100 | < 5 | 13.8 | 2.25 | < 10 |  |  |  |  | 1
| 1 | 3 | < 20 | < 15 | < 20 | < 5 | < 100 | < 100 | < 100 | < 5 | 13.5 | 2 | < 10 |  |  |  |  | 16
| 1 | 4 | < 20 | < 15 | < 20 | < 5 | 100 | < 100 | < 100 | < 5 | 11 | 3.4 | < 10 |  |  |  |  | < 100
| 1 | 5 | < 20 | < 15 | < 20 | < 5 | 100 | < 100 | < 100 | < 5 | 12.3 | 2 | < 10 |  |  |  |  | 21
| 1 | 6 | < 20 | < 15 | < 20 | < 5 | 100 | < 100 | < 100 | < 5 | 16 | 1.5 | < 10
| Mean |  |  |  |  | 5 | 100 |  |  |  | 13.2 | 2.0667
| Median |  |  |  |  | 5 | 100 |  |  |  | 13.05 | 2
| Std Dev. | [5 empty] | 0 |  |  |  | 1.6912 | 0.74944
| Rel.Std.Dev. | [5 empty] | 0% |  |  |  | 12.81% | 36.26%
PDM3
| Z-Score (Absolute) |  | 0 | 0.24 | 0 | 0.67 | 7.77 | 3.81 | 3.81 | 0.71 | 0.3 | 0.76 | 0.48
| Analytical results for B in OREAS 196 (Indicative Value 40.2 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 17 | 20 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | [9 empty] | 1
1
| 1 | 1 | 71 | < 50 | < 50 | < 10 | 33.8 | < 10 | [9 empty] | 1
| 1 | 2 | 64 | < 50 | < 50 | 10 | 46.6 | < 10 | [9 empty] | 16
| 1 | 3 | 53 | < 50 | < 50 | 10 | 24.3333 | < 10 | [9 empty] | 16
| 1 | 4 | 71 | < 50 | < 50 | < 10 | 55.6 | < 10 | [9 empty] | 40.2074
| 1 | 5 | 50 | < 50 | < 50 | < 10 | 64.4 | < 10 | [9 empty] | 22
| 1 | 6 | 83 | < 50 | < 50 | 10 | 62 | < 10
| Mean |  | 65.3333 |  |  | 10 | 47.7889
| Median |  | 67.5 |  |  | 10 | 51.1
| Std Dev. |  | 12.372 |  |  | 0 | 16.0385
| Rel.Std.Dev. |  | 18.94% |  |  | 0% | 33.56%
| PDM3 |  | 62.49% |  |  | -75.13% | 18.86%
| Z-Score (Absolute) |  | 1.45 | 0 | 0 | 0.63 | 0.82 | 0.72
| Analytical results for Ba in OREAS 196 (Certified Value 34.9 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | BF*MS | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 1
2
| 1 | 1 | 34 | 33.1005 | 39.8547 | 36 | 39.8924 | 33 | 35 | 40 | 26.6 | 37.8 | 38 |  |  |  |  | 1
| 1 | 2 | 31 | 34.3816 | 31.893 | 36 | 42.5236 | 33 | 35 | 40 | 23.2 | 31.6 | 36 |  |  |  |  | #N/A
| 1 | 3 | 32 | 34.0096 | 36.0932 | 38 | 39.5828 | 34 | 36 | 40 | 24.7 | 33.4 | 35 |  |  |  |  | 16
| 1 | 4 | 32 | 33.4607 | 41.4783 | 38 | 39.0368 | 34 | 34 | 40 | 22 | 30 | 34 |  |  |  |  | 34.9416
| 1 | 5 | 32 | 33.567 | 36.8567 | 35 | 38.3468 | 35 | 35 | 40 | 23.7 | 28.6 | 38 |  |  |  |  | 28
| 1 | 6 | 31 | 33.1649 | 34.7412 | 35 | 39.0644 | 34 | 35 | 40 | 22.8 | 26.4 | 34
| Mean |  | 32 | 33.6141 | 36.8195 | 36.3333 | 39.7411 | 33.8333 | 35 | 40 | 23.8333 | 31.3 | 35.8333
| Median |  | 32 | 33.5139 | 36.4749 | 36 | 39.3236 | 34 | 35 | 40 | 23.45 | 30.8 | 35.5
| Std Dev. |  | 1.0954 | 0.49734 | 3.4651 | 1.3663 | 1.4621 | 0.75277 | 0.63246 | 0 | 1.6281 | 3.9935 | 1.8348
| Rel.Std.Dev. |  | 3.423% | 1.48% | 9.411% | 3.76% | 3.679% | 2.225% | 1.807% | 0% | 6.831% | 12.76% | 5.121%
| PDM3 |  | -8.419% | -3.799% | 5.374% | 3.983% | 13.74% | -3.172% | 0.167% | 14.48% | -31.79% | -10.42% | 2.552%
| Z-Score (Absolute) |  | 0.75 | 0.25 | 0.75 | 0.6 | 1.66 | 0.18 | 0.18 | NA | 3.3 | 0.97 | 0.44
Results from laboratory 14 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 196 (Certified Value < 1 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 12 | 15 | 20 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | [7 empty] | 2
2
| 1 | 1 | 11.7 | < 1 | < 0.5 | < 1 | < 3 | < 5 | < 2 | < 0.5 | [7 empty] | 1
| 1 | 2 | 11.5 | < 1 |  | < 1 | < 3 | < 5 | < 2 | < 0.5 | [7 empty] | #N/A
| 1 | 3 | 10.8 | < 1 | < 0.5 | < 1 | < 3 | < 5 | < 2 | < 0.5 | [7 empty] | 16
| 1 | 4 | 11.4 | < 1 | < 0.5 | < 1 | < 3 | < 5 | < 2 | < 0.5 | [7 empty] | < 1
| 1 | 5 | 12.7 | < 1 | < 0.5 | < 1 | < 3 | < 5 | < 2 | < 0.5 | [7 empty] | 29
| 1 | 6 | 11.9 | < 1 | < 0.5 | < 1 | < 3 | < 5 | < 2 | < 0.5
| Mean |  | 11.6667
| Median |  | 11.6
| Std Dev. |  | 0.62823
| Rel.Std.Dev. |  | 5.385%
PDM3
| Z-Score (Absolute) |  | 14.72 | 0.34 | 0.67 | 0.34 | 1.01 | 2.36 | 0.34 | 0.67
| Analytical results for Bi in OREAS 196 (Certified Value < 0.1 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 12 | 14 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 2 | < 0.1 | < 3 | < 5 | 1.3 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 2 | < 0.1 | < 3 | < 5 |  | < 0.1 | [7 empty] | 3
| 1 | 3 | < 0.1 | < 0.1 | < 2 | < 0.1 | < 3 | < 5 | 0.25 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 2 | < 0.1 | < 3 | < 5 |  | < 0.1 | [7 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 2 | < 0.1 | < 3 | < 5 | 0.1 | < 0.1 | [7 empty] | 30
| 1 | 6 | < 0.1 | < 0.1 | < 2 | < 0.1 | < 3 | < 5 |  | < 0.1
| Mean | [7 empty] | 0.55
| Median | [7 empty] | 0.25
| Std Dev. | [7 empty] | 0.65383
| Rel.Std.Dev. | [7 empty] | 118.9%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 1.89 | 0.67 | 3.24 | 5.93 | 0.67 | 0.67
| Analytical results for Ca in OREAS 196 (Certified Value 0.219 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.2 | 0.21788 | 0.10653 | 0.22 | 0.172 | 0.229 | 0.222 | 0.25959 | 0.4 | 0.3 | 0.21 | 0.23 | 0.172 | 0.2 |  | 1
| 1 | 2 | 0.2 | 0.21967 | 0.13095 | 0.22 | 0.172 | 0.222 | 0.207 | 0.25617 | 0.4 | 0.2 | 0.22 | 0.23 | 0.268 | 0.2 |  | #N/A
| 1 | 3 | 0.2 | 0.21706 | 0.10272 | 0.17 | 0.172 | 0.222 | 0.222 | 0.2598 | 0.4 | 0.2 | 0.22 | 0.24 | 0.15 | 0.2 |  | 16
| 1 | 4 | 0.2 | 0.21703 | 0.11686 | 0.24 | 0.172 | 0.243 | 0.243 | 0.25533 | 0.4 | 0.2 | 0.21 | 0.23 | 0.258 | 0.2 |  | 0.21895
| 1 | 5 | 0.2 | 0.21413 | 0.12413 | 0.19 | 0.157 | 0.222 | 0.25 | 0.25645 | 0.4 | 0.2 | 0.21 | 0.23 | 0.198 | 0.2 |  | 31
| 1 | 6 | 0.2 | 0.21637 | 0.10329 | 0.21 | 0.172 | 0.222 | 0.293 | 0.25832 | 0.4 | 0.2 | 0.21 | 0.23 | 0.2 | 0.2
| Mean |  | 0.2 | 0.21702 | 0.11408 | 0.20833 | 0.1695 | 0.22667 | 0.2395 | 0.25761 | 0.4 | 0.21667 | 0.21333 | 0.23167 | 0.20767 | 0.2
| Median |  | 0.2 | 0.21704 | 0.11169 | 0.215 | 0.172 | 0.222 | 0.2325 | 0.25739 | 0.4 | 0.2 | 0.21 | 0.23 | 0.199 | 0.2
| Std Dev. |  | 3.0405e-17 | 0.0018208 | 0.011797 | 0.024833 | 0.0061237 | 0.0084774 | 0.030507 | 0.0018893 | 6.0809e-17 | 0.040825 | 0.005164 | 0.0040825 | 0.04674 | 3.0405e-17
| Rel.Std.Dev. |  | 1.52e-14% | 0.839% | 10.34% | 11.92% | 3.613% | 3.74% | 12.74% | 0.7334% | 1.52e-14% | 18.84% | 2.421% | 1.762% | 22.51% | 1.52e-14%
| PDM3 |  | -8.654% | -0.8788% | -47.9% | -4.847% | -22.58% | 3.526% | 9.387% | 17.66% | 82.69% | -1.041% | -2.564% | 5.81% | -5.152% | -8.654%
| Z-Score (Absolute) |  | NA | 0.09 | 4.93 | 0.33 | 2.25 | 0.54 | 1.18 | 2.07 | NA | NA | 0.09 | 0.8 | 0.37 | NA
Results from laboratories 2, 11, 12 and 20 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 196 (Certified Value < 10 ppm) | [16 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS | [7 empty] | 1
1
| 1 | 1 | < 10 | < 10 | < 2 | 12.869 | 1.2 | < 1 | 0.675 | 1 | [7 empty] | 1
| 1 | 2 | < 10 | < 10 | < 2 | 13.0645 | 1.3 | < 1 | 1.35 | 0.9 | [7 empty] | 4
| 1 | 3 | < 10 | < 10 | < 2 | 12.6045 | 1.3 | < 1 | 0.825 | 1 | [7 empty] | 16
| 1 | 4 | < 10 | < 10 | < 2 | 12.731 | 1.3 | < 1 | 2.375 | 1 | [7 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 2 | 13.421 | 1.3 | < 1 | 2.35 | 1.1 | [7 empty] | 32
| 1 | 6 | < 10 | < 10 | < 2 | 11.995 | 1.3 | < 1 | 1.425 | 1.2
| Mean |  |  |  |  | 12.7808 | 1.2833 |  | 1.5 | 1.0333
| Median |  |  |  |  | 12.8 | 1.3 |  | 1.3875 | 1
| Std Dev. |  |  |  |  | 0.47941 | 0.040825 |  | 0.72835 | 0.10328
| Rel.Std.Dev. |  |  |  |  | 3.751% | 3.181% |  | 48.56% | 9.995%
PDM3
| Z-Score (Absolute) |  | 3.79 | 3.79 | 0.41 | 11.97 | 0.11 | 0.94 | 0.11 | 0.38
| Analytical results for Ce in OREAS 196 (Certified Value 2.88 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 05 | 12 | 14 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
3
| 1 | 1 | 6 | 2.6 | 2.6 | 2.7 | 3 | < 0.3 | 2.66 | 3.6 | [7 empty] | 1
| 1 | 2 | 5 | 2.6 | 2.8 | 2.9 | 3 | < 0.3 | 2.76 | 2.9 | [7 empty] | 5
| 1 | 3 | 4 | 2.7 | 2.7 | 2.7 | 3 | < 0.3 | 3.7 | 3.3 | [7 empty] | 16
| 1 | 4 | 4 | 3 | 2.9 | 2.9 | 3 | < 0.3 | 2.72 | 3.2 | [7 empty] | 2.8756
| 1 | 5 | 5 | 2.6 | 3.2 | 2.8 | 3 | < 0.3 | 2.7 | 3.2 | [7 empty] | 33
| 1 | 6 | 6 | 2.7 | 2.9 | 2.9 | 3 | < 0.3 | 2.8 | 3.5
| Mean |  | 5 | 2.7 | 2.85 | 2.8167 | 3 |  | 2.89 | 3.2833
| Median |  | 5 | 2.65 | 2.85 | 2.85 | 3 |  | 2.74 | 3.25
| Std Dev. |  | 0.89443 | 0.15492 | 0.20736 | 0.098319 | 0 |  | 0.39975 | 0.24833
| Rel.Std.Dev. |  | 17.89% | 5.738% | 7.276% | 3.491% | 0% |  | 13.83% | 7.563%
| PDM3 |  | 73.88% | -6.107% | -0.8902% | -2.049% | 4.326% |  | 0.5008% | 14.18%
| Z-Score (Absolute) |  | NA | 0.95 | 0.12 | 0.12 | NA | 19.05 | 0.4 | 3.19
Results from laboratory 1 and 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 196 (Certified Value 574 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | 0
0
| 1 | 1 | 559 | 594 | 571.0967 | 587.523 | 559 | 550 | 550 | 530 | 603.1356 | 574 | 621 | 550 | 604.9 | 661 | 573 | 1
| 1 | 2 | 569 | 590 | 569.9771 | 552.707 | 547 | 550 | 520 | 540 | 608.7982 | 588 | 636 | 550 | 601.8 | 599.6 | 563 | 21
| 1 | 3 | 588 | 595 | 565.9417 | 558.0038 | 545 | 540 | 570 | 540 | 603.6869 | 588 | 629 | 560 | 605 | 567.4 | 548 | 16
| 1 | 4 | 548 | 603 | 578.1731 | 578.2868 | 548 | 520 | 540 | 530 | 606.5779 | 598 | 639 | 560 | 603.1 | 579.2 | 559 | 574.1102
| 1 | 5 | 566 | 599 | 584.982 | 583.518 | 553 | 540 | 550 | 540 | 603.1717 | 596 | 637 | 550 | 603.1 | 582.4 | 567 | 34
| 1 | 6 | 564 | 602 | 573.0177 | 568.3009 | 551 | 550 | 550 | 530 | 603.657 | 597 | 624 | 550 | 606.1 | 594.2 | 561
| Mean |  | 565.6667 | 597.1667 | 573.8647 | 571.3899 | 550.5 | 541.6667 | 546.6667 | 535 | 604.8379 | 590.1667 | 631 | 553.3333 | 604 | 597.3 | 561.8333
| Median |  | 565 | 597 | 572.0572 | 573.2938 | 549.5 | 545 | 550 | 535 | 603.6719 | 592 | 632.5 | 550 | 604 | 588.3 | 562
| Std Dev. |  | 13.1859 | 5.0365 | 6.759 | 14.0905 | 5.0498 | 11.6905 | 16.3299 | 5.4772 | 2.3283 | 9.0866 | 7.4565 | 5.164 | 1.5925 | 33.2151 | 8.4004
| Rel.Std.Dev. |  | 2.331% | 0.8434% | 1.178% | 2.466% | 0.9173% | 2.158% | 2.987% | 1.024% | 0.3849% | 1.54% | 1.182% | 0.9332% | 0.2637% | 5.561% | 1.495%
| PDM3 |  | -1.471% | 4.016% | -0.04275% | -0.4738% | -4.112% | -5.651% | -4.78% | -6.812% | 5.352% | 2.797% | 9.909% | -3.619% | 5.206% | 4.039% | -2.138%
| Z-Score (Absolute) |  | 0.16 | 0.7 | 0.07 | 0 | 0.57 | 0.81 | 0.67 | 0.99 | 0.91 | 0.51 | 1.63 | 0.49 | 0.89 | 0.71 | 0.26
| Analytical results for Cr in OREAS 196 (Certified Value 0.9 wt.%) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 17 | 20 |  |  | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES |  |  | 3
3
| 1 | 1 | 0.85 | 0.91097 | 0.99234 | 0.91 | 0.876 | 0.917 | 0.91 | 0.94659 | 0.9321 | 0.9253 | 0.896 | 0.9812 | 0.6305 |  |  | 1
| 1 | 2 | 0.859 | 0.90473 | 0.97124 | 0.899 | 0.903 | 0.883 | 0.903 | 0.94714 | 0.9249 | 0.9167 | 0.892 | 0.8798 | 0.6504 |  |  | #N/A
| 1 | 3 | 0.859 | 0.90571 | 0.99037 | 0.89 | 0.889 | 0.896 | 0.91 | 0.94284 | 0.898 | 0.9045 | 0.874 | 0.6848 | 0.6326 |  |  | 16
| 1 | 4 | 0.873 | 0.90783 | 0.97259 | 0.911 | 0.876 | 0.91 | 0.903 | 0.94351 | 0.9171 | 0.9216 | 0.896 | 0.8682 | 0.636 |  |  | 0.90003
| 1 | 5 | 0.873 | 0.92248 | 0.98164 | 0.899 | 0.876 | 0.896 | 0.91 | 0.94502 | 0.9162 | 0.9299 | 0.864 | 0.84 | 0.6294 |  |  | 35
| 1 | 6 | 0.858 | 0.90281 | 0.97099 | 0.914 | 0.876 | 0.91 | 0.876 | 0.94131 | 0.9197 | 0.9235 | 0.87 | 0.8008 | 0.624
| Mean |  | 0.862 | 0.90909 | 0.97986 | 0.90383 | 0.88267 | 0.902 | 0.902 | 0.9444 | 0.918 | 0.92025 | 0.882 | 0.84247 | 0.63382
| Median |  | 0.859 | 0.90677 | 0.97711 | 0.9045 | 0.876 | 0.903 | 0.9065 | 0.94427 | 0.9184 | 0.92255 | 0.883 | 0.8541 | 0.63155
| Std Dev. |  | 0.0091652 | 0.0071339 | 0.0097485 | 0.0092826 | 0.011237 | 0.012538 | 0.013191 | 0.0022566 | 0.011431 | 0.0088507 | 0.014311 | 0.097881 | 0.0090347
| Rel.Std.Dev. |  | 1.063% | 0.7847% | 0.9949% | 1.027% | 1.273% | 1.39% | 1.462% | 0.2389% | 1.245% | 0.9618% | 1.623% | 11.62% | 1.425%
| PDM3 |  | -4.225% | 1.007% | 8.87% | 0.4228% | -1.929% | 0.2191% | 0.2191% | 4.93% | 1.997% | 2.247% | -2.003% | -6.396% | -29.58%
| Z-Score (Absolute) |  | 1.39 | 0.25 | 2.71 | 0.06 | 0.67 | 0 | 0 | 1.47 | 0.56 | 0.63 | 0.7 | 2.07 | 9.33
| Analytical results for Cs in OREAS 196 (Certified Value 0.19 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | 14 | 0.23 | < 0.3 | 0.20093 | 0.2 | 0.1847 | 0.2 | 0.1 | < 1 |  | 0.2 |  |  |  |  | 1
| 1 | 2 | 2 | 0.23 | < 0.3 | 0.18137 | 0.8 | 0.3018 |  | 0.1 | < 1 |  | < 0.2 |  |  |  |  | 7
| 1 | 3 | < 2 | 0.24 | < 0.3 | 0.17164 | 1.5 | 0.1885 | 0.2 | < 0.1 | < 1 | 0.2 | 0.2 |  |  |  |  | 16
| 1 | 4 | 2 | 0.22 | < 0.3 | 0.18301 | 1.1 | 0.1561 | 0.2 | 0.2 | < 1 |  | < 0.2 |  |  |  |  | 0.18819
| 1 | 5 | 9 | 0.21 | < 0.3 | 0.1911 | 0.6 | 0.1837 | 0.1 | 0.1 | < 1 |  | < 0.2 |  |  |  |  | 36
| 1 | 6 | 5 | 0.19 | < 0.3 | 0.2031 | 1.2 | 0.1642 | 0.1 | 0.2 | < 1 |  | 0.3
| Mean |  | 6.4 | 0.22 |  | 0.18853 | 0.9 | 0.1965 | 0.16 | 0.14 |  | 0.2 | 0.23333
| Median |  | 5 | 0.225 |  | 0.18706 | 0.95 | 0.1842 | 0.2 | 0.1 |  | 0.2 | 0.2
| Std Dev. |  | 5.1284 | 0.017889 |  | 0.012162 | 0.46476 | 0.053161 | 0.054772 | 0.054772 |  |  | 0.057735
| Rel.Std.Dev. |  | 80.13% | 8.131% |  | 6.451% | 51.64% | 27.05% | 34.23% | 39.12% |  |  | 24.74%
| PDM3 |  | 3301% | 16.91% |  | 0.1816% | 378.3% | 4.418% | -14.98% | -25.61% |  | 6.278% | 23.99%
| Z-Score (Absolute) |  | NA | 0.62 | 0.96 | 0.09 | 15.9 | 0.09 | 0.73 | 1.52 | 6.91 | 0.17 | 0.58
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 196 (Certified Value 45 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 08 | 09 | 11 | 12 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 1
2
| 1 | 1 | 40 | 47.9769 | 49 | 30 | 40 | 30 | 21 | 45 | 42.4 | 39.8 | 42 |  |  |  |  | 1
| 1 | 2 | 40 | 45.1847 | 49 | 30 | 40 | 30 |  | 44 | 44.5 | 44 | 46 |  |  |  |  | #N/A
| 1 | 3 | 40 | 47.3818 | 42 | 40 | 40 | 20 | 23 | 44 | 42.13 | 38.4 | 45 |  |  |  |  | 16
| 1 | 4 | 40 | 47.7523 | 44 | 30 | 30 | 20 | 23 | 45 | 43.88 | 46.6 | 47 |  |  |  |  | 45.0113
| 1 | 5 | 40 | 48.414 | 45 | 40 | 50 | 30 | 26 | 46 | 41.62 | 54.4 | 42 |  |  |  |  | 37
| 1 | 6 | 40 | 45.3759 | 44 | 40 | 40 | 20 | 30 | 46 | 40.59 | 52 | 43
| Mean |  | 40 | 47.0143 | 45.5 | 35 | 40 | 25 | 24.6 | 45 | 42.52 | 45.8667 | 44.1667
| Median |  | 40 | 47.567 | 44.5 | 35 | 40 | 25 | 23 | 45 | 42.265 | 45.3 | 44
| Std Dev. |  | 0 | 1.3855 | 2.881 | 5.4772 | 6.3246 | 5.4772 | 3.5071 | 0.89443 | 1.4472 | 6.4345 | 2.137
| Rel.Std.Dev. |  | 0% | 2.947% | 6.332% | 15.65% | 15.81% | 21.91% | 14.26% | 1.988% | 3.403% | 14.03% | 4.838%
| PDM3 |  | -11.13% | 4.45% | 1.086% | -22.24% | -11.13% | -44.46% | -45.35% | -0.02503% | -5.535% | 1.9% | -1.876%
| Z-Score (Absolute) |  | NA | 1.57 | 0.39 | NA | NA | NA | 15.87 | 0 | 1.93 | 0.67 | 0.65
Results from laboratories 2, 7, 8 and 9 were removed due to their 10 ppm reading resolution.
| Analytical results for Dy in OREAS 196 (Certified Value 1.58 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | < 2 | 1.7 | 1.76 | 1.5 | 1.2594 | 1.5 | 1.5 | 2 | 1.44 | 1.85 | [5 empty] | 1
| 1 | 2 | < 2 | 1.7 | 1.5572 | 1.5 | 1.2647 | 1.51 | 2 | 2 | 1.458 | 1.6 | [5 empty] | #N/A
| 1 | 3 | < 2 | 1.8 | 1.5406 | 1.4 | 1.2398 | 1.48 | 1.5 | 1.9 | 1.348 | 1.71 | [5 empty] | 16
| 1 | 4 | < 2 | 1.7 | 1.5596 | 1.6 | 1.143 | 1.55 | 1.5 | 1.9 | 1.39 | 1.68 | [5 empty] | 1.5812
| 1 | 5 | < 2 | 1.7 | 1.6788 | 1.5 | 1.1933 | 1.55 | 1.5 | 1.9 | 1.408 | 1.69 | [5 empty] | 38
| 1 | 6 | < 2 | 1.7 | 1.5652 | 1.7 | 1.4314 | 1.58 | 1.5 | 1.9 | 1.426 | 1.92
| Mean |  |  | 1.7167 | 1.6102 | 1.5333 | 1.2553 | 1.5283 | 1.5833 | 1.9333 | 1.4117 | 1.7417
| Median |  |  | 1.7 | 1.5624 | 1.5 | 1.2496 | 1.53 | 1.5 | 1.9 | 1.417 | 1.7
| Std Dev. |  |  | 0.040825 | 0.088756 | 0.10328 | 0.097793 | 0.037639 | 0.20412 | 0.05164 | 0.039221 | 0.11923
| Rel.Std.Dev. |  |  | 2.378% | 5.512% | 6.736% | 7.791% | 2.463% | 12.89% | 2.671% | 2.778% | 6.846%
| PDM3 |  |  | 8.57% | 1.839% | -3.025% | -20.61% | -3.341% | 0.137% | 22.27% | -10.72% | 10.15%
| Z-Score (Absolute) |  | 2.47 | 0.7 | 0.23 | 0.11 | 1.34 | 0.13 | 0.11 | 1.66 | 0.65 | 0.81
| Analytical results for Er in OREAS 196 (Certified Value 0.94 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
3
| 1 | 1 | < 2 | 1 | 0.87663 | 0.9 | 0.9285 | 0.95 | < 1 | 1.3 | 0.988 | 1.05 | [5 empty] | 1
| 1 | 2 | < 2 | 1 | 0.94698 | 0.9 | 0.9524 | 0.92 | < 1 | 1.4 | 0.952 | 0.97 | [5 empty] | #N/A
| 1 | 3 | < 2 | 1 | 0.93759 | 0.9 | 0.9396 | 0.86 | < 1 | 1.3 | 0.826 | 0.99 | [5 empty] | 16
| 1 | 4 | < 2 | 1 | 0.86319 | 1 | 0.9506 | 0.9 | < 1 | 1.4 | 0.864 | 0.93 | [5 empty] | 0.93545
| 1 | 5 | < 2 | 1 | 0.88741 | 0.9 | 0.9681 | 0.9 | < 1 | 1.3 | 0.832 | 1 | [5 empty] | 39
| 1 | 6 | < 2 | 1 | 0.99298 | 1 | 0.9635 | 0.89 | < 1 | 1.3 | 0.944 | 1.09
| Mean |  |  | 1 | 0.91746 | 0.93333 | 0.95045 | 0.90333 |  | 1.3333 | 0.901 | 1.005
| Median |  |  | 1 | 0.9125 | 0.9 | 0.9515 | 0.9 |  | 1.3 | 0.904 | 0.995
| Std Dev. |  |  | 0 | 0.049989 | 0.05164 | 0.014732 | 0.030111 |  | 0.05164 | 0.068955 | 0.057184
| Rel.Std.Dev. |  |  | 0% | 5.449% | 5.533% | 1.55% | 3.333% |  | 3.873% | 7.653% | 5.69%
| PDM3 |  |  | 6.9% | -1.923% | -0.2262% | 1.604% | -3.433% |  | 42.53% | -3.683% | 7.435%
| Z-Score (Absolute) |  | 1.69 | NA | 0 | NA | 0.67 | 0.29 | 8.53 | NA | 0.34 | 1.79
Results from laboratories 2, 5 and 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 196 (Certified Value 0.51 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | < 2 | 0.5 | 0.42429 | 0.5 | 0.5839 | 0.49 | < 0.5 | < 0.5 | 0.68 | 0.62 | [5 empty] | 1
| 1 | 2 | < 2 | 0.5 | 0.42851 | 0.5 | 0.5227 | 0.5 | < 0.5 | < 0.5 | 0.498 | 0.59 | [5 empty] | #N/A
| 1 | 3 | < 2 | 0.5 | 0.44137 | 0.6 | 0.5505 | 0.5 | < 0.5 | < 0.5 | 0.46 | 0.57 | [5 empty] | 16
| 1 | 4 | < 2 | 0.5 | 0.42101 | 0.5 | 0.4949 | 0.49 | < 0.5 | < 0.5 | 0.502 | 0.6 | [5 empty] | 0.513
| 1 | 5 | < 2 | 0.5 | 0.40574 | 0.5 | 0.481 | 0.52 | < 0.5 | < 0.5 | 0.5 | 0.61 | [5 empty] | 40
| 1 | 6 | < 2 | 0.5 | 0.38854 | 0.5 | 0.6102 | 0.53 | < 0.5 | < 0.5 | 0.496 | 0.67
| Mean |  |  | 0.5 | 0.41824 | 0.51667 | 0.54053 | 0.505 |  |  | 0.52267 | 0.61
| Median |  |  | 0.5 | 0.42265 | 0.5 | 0.5366 | 0.5 |  |  | 0.499 | 0.605
| Std Dev. |  |  | 0 | 0.018564 | 0.040825 | 0.050553 | 0.016432 |  |  | 0.078666 | 0.034059
| Rel.Std.Dev. |  |  | 0% | 4.439% | 7.902% | 9.352% | 3.254% |  |  | 15.05% | 5.583%
| PDM3 |  |  | -2.533% | -18.47% | 0.7157% | 5.368% | -1.559% |  |  | 1.885% | 18.91%
| Z-Score (Absolute) |  | 3.42 | NA | 0.67 | NA | 0.19 | 0.06 | 1.86 | 1.86 | 0.06 | 0.68
Results from laboratories 2 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 196 (Certified Value 18.74 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 | wt.%
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | 18.1496 | 18.4 | 18.5327 | 19.0784 | 18.8 | 18.535 | 18.605 | 18.2 | 19.2542 | 18.87 | 18.82 | 18.8 | 18.82 | 20.214 | 18.16 | 1
| 1 | 2 | 18.6235 | 18.6 | 18.4248 | 19.201 | 18.7 | 18.605 | 18.116 | 17.9 | 19.1706 | 18.78 | 18.67 | 18.5 | 18.73 | 19.9762 | 19.48 | 22
| 1 | 3 | 18.6306 | 18.7 | 18.3346 | 19.1945 | 18.6 | 18.885 | 18.605 | 17.85 | 19.2979 | 18.26 | 18.36 | 18.4 | 18.76 | 18.8571 | 18.87 | 16
| 1 | 4 | 18.4509 | 19 | 18.8406 | 19.1267 | 18.7 | 18.326 | 18.675 | 17.8 | 19.1223 | 18.64 | 18.88 | 18.4 | 18.73 | 19.7384 | 19.7 | 18.7429
| 1 | 5 | 18.73 | 19 | 18.8548 | 19.2147 | 18.7 | 18.815 | 18.535 | 18.1 | 19.0425 | 18.6 | 18.88 | 18.3 | 19.02 | 18.7172 | 18.46 | 41
| 1 | 6 | 17.8157 | 18.7 | 18.6872 | 19.0546 | 18.6 | 19.025 | 18.605 | 17.95 | 19.0193 | 18.68 | 18.73 | 18.4 | 18.75 | 18.1157 | 19.13
| Mean |  | 18.4001 | 18.7333 | 18.6124 | 19.145 | 18.6833 | 18.6985 | 18.5235 | 17.9667 | 19.1511 | 18.6383 | 18.7233 | 18.4667 | 18.8017 | 19.2698 | 18.9667
| Median |  | 18.5372 | 18.7 | 18.61 | 19.1606 | 18.7 | 18.71 | 18.605 | 17.925 | 19.1464 | 18.66 | 18.775 | 18.4 | 18.755 | 19.2978 | 19
| Std Dev. |  | 0.35177 | 0.23381 | 0.21687 | 0.068395 | 0.075277 | 0.25659 | 0.20448 | 0.15384 | 0.11188 | 0.2098 | 0.19664 | 0.17512 | 0.11197 | 0.82678 | 0.59072
| Rel.Std.Dev. |  | 1.912% | 1.248% | 1.165% | 0.3572% | 0.4029% | 1.372% | 1.104% | 0.8563% | 0.5842% | 1.126% | 1.05% | 0.9483% | 0.5955% | 4.291% | 3.115%
| PDM3 |  | -1.829% | -0.0513% | -0.6963% | 2.145% | -0.3181% | -0.2372% | -1.171% | -4.142% | 2.178% | -0.5582% | -0.1047% | -1.474% | 0.3133% | 2.811% | 1.194%
| Z-Score (Absolute) |  | 1.15 | 0.13 | 0.33 | 1.72 | 0.06 | 0 | 0.67 | 2.82 | 1.75 | 0.23 | 0.1 | 0.89 | 0.4 | 2.2 | 1.03
| Analytical results for Ga in OREAS 196 (Certified Value 4.22 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | 4 | 5.0426 | 4.538 | 3.6 | 2.9219 | 4 | 5 | 5 | < 5 | 4.8 | 4 |  |  |  |  | 1
| 1 | 2 | 4 | 5.1632 | 4.2327 | 2.7 | 2.87 | 4 | 4 | 5 | < 5 | 5.6 | 4 |  |  |  |  | 8
| 1 | 3 | 4 | 5.2217 | 4.119 | 4.1 | 2.9271 | 4 | 4 | 6 | < 5 | 7.8 | 4 |  |  |  |  | 16
| 1 | 4 | 4 | 4.8823 | 4.1849 | 4.1 | 2.8459 | 4 | 4 | 5 | < 5 | 5.2 | 4 |  |  |  |  | 4.2168
| 1 | 5 | 4 | 5.3349 | 4.233 | 4.1 | 2.9568 | 4 | 4 | 6 | < 5 | 5.4 | 4 |  |  |  |  | 42
| 1 | 6 | 4 | 4.922 | 4.4943 | 3.5 | 2.7927 | 4 | 4 | 5 | < 5 | 4.6 | 5
| Mean |  | 4 | 5.0944 | 4.3003 | 3.6833 | 2.8857 | 4 | 4.1667 | 5.3333 |  | 5.5667 | 4.1667
| Median |  | 4 | 5.1029 | 4.2328 | 3.85 | 2.896 | 4 | 4 | 5 |  | 5.3 | 4
| Std Dev. |  | 0 | 0.17673 | 0.17288 | 0.55287 | 0.06091 | 0 | 0.40825 | 0.5164 |  | 1.1553 | 0.40825
| Rel.Std.Dev. |  | 0% | 3.469% | 4.02% | 15.01% | 2.111% | 0% | 9.798% | 9.682% |  | 20.75% | 9.798%
| PDM3 |  | -5.141% | 20.81% | 1.981% | -12.65% | -31.56% | -5.141% | -1.188% | 26.48% |  | 32.01% | -1.188%
| Z-Score (Absolute) |  | NA | 0.67 | 0.19 | 0.19 | 0.68 | NA | NA | NA | 0.91 | 0.96 | NA
Results from laboratories 2, 11, 12, 14 and 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 196 (Certified Value 1.73 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | < 2 | 1.7 | 1.9 | 1.5957 | 1.72 | 2 | < 0.5 | 1.858 | 1.94 | [6 empty] | 1
| 1 | 2 | < 2 | 1.8 | 1.6 | 1.6134 | 1.71 | 2 | < 0.5 | 1.784 | 1.71 | [6 empty] | #N/A
| 1 | 3 | < 2 | 1.6 | 1.5 | 1.6445 | 1.65 | 2 | < 0.5 | 1.806 | 1.83 | [6 empty] | 16
| 1 | 4 | < 2 | 1.8 | 1.6 | 1.5602 | 1.7 | 2 | < 0.5 | 1.744 | 1.88 | [6 empty] | 1.7288
| 1 | 5 | < 2 | 1.7 | 1.6 | 1.6111 | 1.65 | 2 | < 0.5 | 1.74 | 1.89 | [6 empty] | 43
| 1 | 6 | < 2 | 1.7 | 1.8 | 1.5744 | 1.87 | 2 | < 0.5 | 1.766 | 2.09
| Mean |  |  | 1.7167 | 1.6667 | 1.5999 | 1.7167 | 2 |  | 1.783 | 1.89
| Median |  |  | 1.7 | 1.6 | 1.6034 | 1.705 | 2 |  | 1.775 | 1.885
| Std Dev. |  |  | 0.075277 | 0.15055 | 0.030118 | 0.080911 | 0 |  | 0.044303 | 0.12538
| Rel.Std.Dev. |  |  | 4.385% | 9.033% | 1.883% | 4.713% | 0% |  | 2.485% | 6.634%
| PDM3 |  |  | -0.7026% | -3.595% | -7.458% | -0.7026% | 15.69% |  | 3.134% | 9.324%
| Z-Score (Absolute) |  | 5.09 | 0.18 | 0.18 | 0.68 | 0.18 | NA | 10.62 | 0.67 | 1.46
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 196 (Certified Value 2.4 ppm) | [16 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 10 | 11 | 12 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 3 | 2.567 | 1.8 | 14.488 | 2 | 3 | 2 | [8 empty] | 1
| 1 | 2 | 3 | 2.4298 | 3 | 14.632 | 2 | 2 | 2 | [8 empty] | 9
| 1 | 3 | 3 | 2.2718 | 2.8 | 14.616 | 2 | 2 | 2 | [8 empty] | 16
| 1 | 4 | 3 | 2.0976 | 2.5 | 14.881 | 2 | 2 | 2 | [8 empty] | 2.4027
| 1 | 5 | 3 | 2.9618 | 2.5 | 14.467 | 2 | 3 | 2 | [8 empty] | 44
| 1 | 6 | 3 | 2.8689 | 2.7 | 14.894 | 2 | 2 | 2
| Mean |  | 3 | 2.5328 | 2.55 | 14.663 | 2 | 2.3333 | 2
| Median |  | 3 | 2.4984 | 2.6 | 14.624 | 2 | 2 | 2
| Std Dev. |  | 0 | 0.33653 | 0.41352 | 0.18606 | 0 | 0.5164 | 0
| Rel.Std.Dev. |  | 0% | 13.29% | 16.22% | 1.269% | 0% | 22.13% | 0%
| PDM3 |  | 24.86% | 5.416% | 6.131% | 510.3% | -16.76% | -2.887% | -16.76%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.02 | 17.51 | 0.77 | 0.29 | 0.77
| Analytical results for Hf in OREAS 196 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 05 | 12 | 14 | 15 | [10 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | [10 empty] | 2
2
| 1 | 1 | 0.31697 | < 10 | < 1 | < 1 | < 0.5 | [10 empty] | 1
| 1 | 2 | 0.39807 | < 10 | < 1 | < 1 | < 0.5 | [10 empty] | 17
| 1 | 3 | 0.27665 | < 10 | < 1 | < 1 | < 0.5 | [10 empty] | 16
| 1 | 4 | 0.28503 | < 10 | < 1 | < 1 | < 0.5 | [10 empty] | < 1
| 1 | 5 | 0.29274 | < 10 | < 1 | < 1 | < 0.5 | [10 empty] | 23
| 1 | 6 | 0.28084 | < 10 | < 1 | < 1 | < 0.5
| Mean |  | 0.30838
| Median |  | 0.28889
| Std Dev. |  | 0.046201
| Rel.Std.Dev. |  | 14.98%
PDM3
| Z-Score (Absolute) |  | 0.67 | 15.84 | 0 | 0 | 0.88
| Analytical results for Hg in OREAS 196 (Indicative Value < 5 ppm) | [16 empty] | Indicative
Hg | Round | Replicate | Lab | [14 empty] | 1
| No | No | 15 | [14 empty] | ppm
|  |  | PF*OES | [14 empty] | 2
2
| 1 | 1 | < 5 | [14 empty] | 1
| 1 | 2 | < 5 | [14 empty] | 18
| 1 | 3 | < 5 | [14 empty] | 16
| 1 | 4 | < 5 | [14 empty] | < 5
| 1 | 5 | < 5 | [14 empty] | 24
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 196 (Certified Value 0.32 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
3
| 1 | 1 | < 2 | 0.3 | 0.31551 | 0.3 | 0.3042 | 0.32 | < 1 | < 0.5 | 0.304 | 0.32 | [5 empty] | 1
| 1 | 2 | < 2 | 0.3 | 0.34658 | 0.3 | 0.3017 | 0.32 | < 1 | < 0.5 | 0.348 | 0.3 | [5 empty] | #N/A
| 1 | 3 | < 2 | 0.4 | 0.32732 | 0.3 | 0.3061 | 0.29 | < 1 | < 0.5 | 0.288 | 0.32 | [5 empty] | 16
| 1 | 4 | < 2 | 0.3 | 0.34037 | 0.3 | 0.2885 | 0.3 | < 1 | < 0.5 | 0.314 | 0.31 | [5 empty] | 0.31643
| 1 | 5 | < 2 | 0.3 | 0.31738 | 0.3 | 0.2803 | 0.3 | < 1 | < 0.5 | 0.312 | 0.33 | [5 empty] | 45
| 1 | 6 | < 2 | 0.3 | 0.33726 | 0.3 | 0.3356 | 0.33 | < 1 | < 0.5 | 0.346 | 0.34
| Mean |  |  | 0.31667 | 0.33074 | 0.3 | 0.30273 | 0.31 |  |  | 0.31867 | 0.32
| Median |  |  | 0.3 | 0.33229 | 0.3 | 0.30295 | 0.31 |  |  | 0.313 | 0.32
| Std Dev. |  |  | 0.040825 | 0.012714 | 0 | 0.018977 | 0.015492 |  |  | 0.023788 | 0.014142
| Rel.Std.Dev. |  |  | 12.89% | 3.844% | 0% | 6.269% | 4.997% |  |  | 7.465% | 4.419%
| PDM3 |  |  | 0.07571% | 4.522% | -5.191% | -4.328% | -2.031% |  |  | 0.7078% | 1.129%
| Z-Score (Absolute) |  | 32.78 | NA | 0.55 | NA | 0.8 | 0.45 | 8.7 | 3.34 | 0.03 | 0.03
Results from laboratories 2 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 196 (Certified Value < 0.1 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 14 | 17 | 20 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 0.2 | < 0.1 | < 0.2 | < 0.5 | 0.1 | < 0.2 | [7 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.2 | < 0.1 | < 0.2 | < 0.5 | 0.056667 | < 0.2 | [7 empty] | 2
| 1 | 3 | < 0.1 | < 0.1 | < 0.2 | < 0.1 | < 0.2 | < 0.5 | 0.083333 | < 0.2 | [7 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.2 | < 0.1 | < 0.2 | < 0.5 | 0.06 | < 0.2 | [7 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 0.2 | < 0.1 | < 0.2 | < 0.5 | 0.06 | < 0.2 | [7 empty] | 46
| 1 | 6 | < 0.1 | < 0.1 | < 0.2 | < 0.1 | < 0.2 | < 0.5 | 0.06 | < 0.2
| Mean | [7 empty] | 0.07
| Median | [7 empty] | 0.06
| Std Dev. | [7 empty] | 0.017638
| Rel.Std.Dev. | [7 empty] | 25.2%
PDM3
| Z-Score (Absolute) |  | 0.94 | 0.94 | 0.4 | 0.94 | 0.4 | 4.45 | 0.4 | 0.4
| Analytical results for K in OREAS 196 (Indicative Value 0.091 wt.%) | [16 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 11 | 12 | 14 | 15 | 17 | 20 |  |  | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 3
3
| 1 | 1 | 0.1465 | 0.14 | < 0.05 | < 0.1 | < 0.05 | 0.066 | < 0.05 | 0.12 | 0.1 | 0.04 | 0.04 | 0.1675 | < 0.1 |  |  | 1
| 1 | 2 | 0.1352 | 0.14 | < 0.05 | < 0.1 | < 0.05 | 0.066 | < 0.05 | 0.11 | < 0.1 | 0.04 | 0.04 | 0.134 | 0.1 |  |  | 19
| 1 | 3 | 0.1136 | 0.14 | < 0.05 | < 0.1 | < 0.05 | 0.058 | 0.091 | 0.13 | 0.1 | 0.04 | 0.04 | 0.21 | 0.1 |  |  | 16
| 1 | 4 | 0.0714 | 0.13 | < 0.05 | < 0.1 | < 0.05 | 0.058 | 0.05 | 0.13 | < 0.1 | 0.04 | 0.04 | 0.146 | < 0.1 |  |  | 0.09114
| 1 | 5 | 0.0918 | 0.13 | < 0.05 | < 0.1 | < 0.05 | 0.066 | 0.066 | 0.12 | < 0.1 | 0.04 | 0.04 | 0.086 | < 0.1 |  |  | 25
| 1 | 6 | 0.1279 | 0.13 | < 0.05 | < 0.1 | < 0.05 | 0.058 | 0.066 | 0.11 | 0.1 | 0.04 | 0.04 | 0.044 | 0.1
| Mean |  | 0.1144 | 0.135 |  |  |  | 0.062 | 0.06825 | 0.12 | 0.1 | 0.04 | 0.04 | 0.13125 | 0.1
| Median |  | 0.12075 | 0.135 |  |  |  | 0.062 | 0.066 | 0.12 | 0.1 | 0.04 | 0.04 | 0.14 | 0.1
| Std Dev. |  | 0.028304 | 0.0054772 |  |  |  | 0.0043818 | 0.016939 | 0.0089443 | 1.6997e-17 | 0 | 0 | 0.058993 | 1.6997e-17
| Rel.Std.Dev. |  | 24.74% | 4.057% |  |  |  | 7.067% | 24.82% | 7.454% | 1.7e-14% | 0% | 0% | 44.95% | 1.7e-14%
| PDM3 |  | 25.52% | 48.12% |  |  |  | -31.97% | -25.12% | 31.67% | 9.721% | -56.11% | -56.11% | 44.01% | 9.721%
| Z-Score (Absolute) |  | 1.58 | 2.2 | 1.13 | 0.37 | 1.13 | 0 | 0.25 | 1.75 | 0.39 | 0.67 | 0.67 | 2.09 | 0.39
| Analytical results for La in OREAS 196 (Certified Value 10.4 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 05 | 12 | 14 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | [7 empty] | 1
2
| 1 | 1 | 13 | 9.4 | 9.7 | 10.3 | 10 | 11 | 10.22 | 11 | [7 empty] | 1
| 1 | 2 | 13 | 10 | 10 | 10.4 | 11 | 11.1 | 10.68 | 11 | [7 empty] | 3
| 1 | 3 | 13 | 9.8 | 9.5 | 10 | 12 | 11.1 | 10.62 | 10 | [7 empty] | 16
| 1 | 4 | 12 | 9.8 | 10.7 | 10.5 | 10 | 10.8 | 10.28 | 10 | [7 empty] | 10.3653
| 1 | 5 | 13 | 10 | 10.4 | 10.4 | 11 | 11.1 | 10.16 | 11 | [7 empty] | 47
| 1 | 6 | 13 | 10 | 10.3 | 10.8 | 11 | 11.2 | 10.7 | < 10
| Mean |  | 12.8333 | 9.8333 | 10.1 | 10.4 | 10.8333 | 11.05 | 10.4433 | 10.6
| Median |  | 13 | 9.9 | 10.15 | 10.4 | 11 | 11.1 | 10.45 | 11
| Std Dev. |  | 0.40825 | 0.23381 | 0.45166 | 0.26077 | 0.75277 | 0.13784 | 0.24897 | 0.54772
| Rel.Std.Dev. |  | 3.181% | 2.378% | 4.472% | 2.507% | 6.949% | 1.247% | 2.384% | 5.167%
| PDM3 |  | 23.81% | -5.132% | -2.56% | 0.3344% | 4.515% | 6.605% | 0.7525% | 2.264%
| Z-Score (Absolute) |  | NA | 1.27 | 0.67 | 0 | NA | 1.46 | 0.1 | NA
Results from laboratories 1, 14 and 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 196 (Indicative Value 3.06 ppm) | [16 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 10 | 11 | 12 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 2
2
| 1 | 1 | 5 | 3 | < 1 | < 15 | 10 | 40 | 2.4483 | 3 | < 10 | 44.2 | < 10 |  |  |  |  | 1
| 1 | 2 | 3 | 3 | < 1 | < 15 | 10 | 50 | 2.775 | 3 | < 10 | 33 | < 10 |  |  |  |  | 14
| 1 | 3 | < 2 | 4 | < 1 | < 15 | 10 | 30 | 2.775 | 2 | < 10 | 32.825 | < 10 |  |  |  |  | 16
| 1 | 4 | < 2 | 3 | < 1 | < 15 | 10 | 50 | 2.5433 | 2 | < 10 | 36 | < 10 |  |  |  |  | 3.065
| 1 | 5 | 4 | 3 | < 1 | < 15 | 10 | 20 | 2.385 | 3 | < 10 | 19.74 | < 10 |  |  |  |  | 20
| 1 | 6 | 3 | 4 | < 1 | < 15 | 10 | 30 | 2.5333 | 7 | < 10 | 31 | < 10
| Mean |  | 3.75 | 3.3333 |  |  | 10 | 36.6667 | 2.5767 | 3.3333 |  | 32.7942
| Median |  | 3.5 | 3 |  |  | 10 | 35 | 2.5383 | 3 |  | 32.9125
| Std Dev. |  | 0.95743 | 0.5164 |  |  | 0 | 12.1106 | 0.16424 | 1.8619 |  | 7.9252
| Rel.Std.Dev. |  | 25.53% | 15.49% |  |  | 0% | 33.03% | 6.374% | 55.86% |  | 24.17%
| PDM3 |  | 22.35% | 8.755% |  |  | 226.3% | 1096% | -15.93% | 8.755% |  | 970%
| Z-Score (Absolute) |  | 0.6 | 0.46 | 1.25 | 0.7 | 1.39 | 8.81 | 0.67 | 0.46 | 0 | 7.73 | 0
| Analytical results for Lu in OREAS 196 (Certified Value 0.12 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 10 | 12 | 14 | 15 | 17 | 20 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | < 2 | 0.08 | 0.15963 | 0.1295 | 0.11 | < 0.5 | < 0.5 | 0.134 | 0.2 | [6 empty] | 1
| 1 | 2 | < 2 | 0.06 | 0.1616 | 0.1187 | 0.12 | < 0.5 | < 0.5 | 0.15 | 0.2 | [6 empty] | #N/A
| 1 | 3 | < 2 | 0.09 | 0.15045 | 0.117 | 0.12 | < 0.5 | < 0.5 | 0.154 | 0.2 | [6 empty] | 16
| 1 | 4 | < 2 | 0.07 | 0.1625 | 0.1162 | 0.12 | < 0.5 | < 0.5 | 0.14 | 0.1 | [6 empty] | 0.12231
| 1 | 5 | < 2 | < 0.05 | 0.14839 | 0.1244 | 0.13 | < 0.5 | < 0.5 | 0.148 | 0.2 | [6 empty] | 48
| 1 | 6 | < 2 | 0.09 | 0.15679 | 0.1192 | 0.13 | < 0.5 | < 0.5 | 0.134 | 0.2
| Mean |  |  | 0.078 | 0.15656 | 0.12083 | 0.12167 |  |  | 0.14333 | 0.18333
| Median |  |  | 0.08 | 0.15821 | 0.11895 | 0.12 |  |  | 0.144 | 0.2
| Std Dev. |  |  | 0.013038 | 0.0059043 | 0.0051228 | 0.0075277 |  |  | 0.0085479 | 0.040825
| Rel.Std.Dev. |  |  | 16.72% | 3.771% | 4.24% | 6.187% |  |  | 5.964% | 22.27%
| PDM3 |  |  | -36.23% | 28% | -1.209% | -0.5281% |  |  | 17.19% | 49.89%
| Z-Score (Absolute) |  | 10.43 | 0.99 | 0.08 | 0.36 | 0.35 | 1.23 | 1.23 | 0.08 | NA
Results from laboratory 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 196 (Certified Value 7.39 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 | wt.%
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | 7.3134 | 7.27 | 7.4046 | 7.5443 | 7.49 | 7.327 | 7.448 | 6.996 | 7.0392 | 7.65 | 7.37 | 7.51 | 7.26 | 7.4 | 7.14 | 1
| 1 | 2 | 7.5085 | 7.34 | 7.4194 | 7.214 | 7.44 | 7.568 | 7.176 | 6.875 | 7.0239 | 7.61 | 7.33 | 7.43 | 7.14 | 7.34 | 7.88 | 21
| 1 | 3 | 7.5814 | 7.39 | 7.2883 | 7.4599 | 7.33 | 7.478 | 7.327 | 6.965 | 7.0237 | 7.43 | 7.19 | 7.4 | 7.28 | 7.51 | 7.47 | 16
| 1 | 4 | 7.447 | 7.22 | 7.4254 | 7.6434 | 7.46 | 7.418 | 7.448 | 6.935 | 7.0316 | 7.56 | 7.4 | 7.45 | 7.26 | 7.43 | 7.5 | 7.3928
| 1 | 5 | 7.5391 | 7.44 | 7.3861 | 7.5074 | 7.48 | 7.357 | 7.357 | 6.965 | 7.0155 | 7.54 | 7.41 | 7.21 | 7.38 | 7.53 | 7.28 | 49
| 1 | 6 | 7.2868 | 7.31 | 7.3898 | 7.1363 | 7.42 | 7.357 | 7.418 | 6.845 | 7.0278 | 7.59 | 7.33 | 7.25 | 7.25 | 7.44 | 7.27
| Mean |  | 7.446 | 7.3283 | 7.3856 | 7.4175 | 7.4367 | 7.4175 | 7.3623 | 6.9302 | 7.027 | 7.5633 | 7.3383 | 7.375 | 7.2617 | 7.4417 | 7.4233
| Median |  | 7.4777 | 7.325 | 7.3972 | 7.4836 | 7.45 | 7.3875 | 7.3875 | 6.95 | 7.0258 | 7.575 | 7.35 | 7.415 | 7.26 | 7.435 | 7.375
| Std Dev. |  | 0.12152 | 0.079854 | 0.050146 | 0.19872 | 0.058195 | 0.091457 | 0.1037 | 0.058448 | 0.0080497 | 0.075807 | 0.080104 | 0.11862 | 0.076529 | 0.070261 | 0.26113
| Rel.Std.Dev. |  | 1.632% | 1.09% | 0.679% | 2.679% | 0.7825% | 1.233% | 1.409% | 0.8434% | 0.1146% | 1.002% | 1.092% | 1.608% | 1.054% | 0.9442% | 3.518%
| PDM3 |  | 0.7203% | -0.8718% | -0.09766% | 0.3349% | 0.5936% | 0.3343% | -0.4119% | -6.258% | -4.949% | 2.307% | -0.7366% | -0.2406% | -1.774% | 0.6612% | 0.4132%
| Z-Score (Absolute) |  | 0.8 | 0.76 | 0 | 0.42 | 0.67 | 0.42 | 0.31 | 6.01 | 4.73 | 2.35 | 0.62 | 0.14 | 1.63 | 0.74 | 0.5
| Analytical results for Mn in OREAS 196 (Certified Value 0.368 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*MS |  | 3
3
| 1 | 1 | 0.32 | 0.37363 | 0.39387 | 0.377 | 0.356 | 0.372 | 0.364 | 0.35715 | 0.3863 | 0.3674 | 0.355 | 0.34449 | 0.41604 | 0.3761 |  | 1
| 1 | 2 | 0.324 | 0.37285 | 0.39957 | 0.373 | 0.372 | 0.364 | 0.356 | 0.35942 | 0.3852 | 0.3643 | 0.355 | 0.34243 | 0.39358 | 0.3624 |  | #N/A
| 1 | 3 | 0.325 | 0.37036 | 0.40514 | 0.367 | 0.364 | 0.372 | 0.364 | 0.35308 | 0.3786 | 0.3595 | 0.35 | 0.34232 | 0.35176 | 0.3662 |  | 16
| 1 | 4 | 0.33 | 0.37284 | 0.40517 | 0.373 | 0.364 | 0.372 | 0.356 | 0.35688 | 0.3839 | 0.3695 | 0.355 | 0.34264 | 0.39064 | 0.3646 |  | 0.36777
| 1 | 5 | 0.33 | 0.37891 | 0.39315 | 0.369 | 0.356 | 0.372 | 0.364 | 0.35413 | 0.3809 | 0.3681 | 0.35 | 0.35224 | 0.36756 | 0.3738 |  | 50
| 1 | 6 | 0.324 | 0.37395 | 0.39935 | 0.368 | 0.364 | 0.372 | 0.356 | 0.3576 | 0.3811 | 0.3658 | 0.355 | 0.34255 | 0.35764 | 0.4177
| Mean |  | 0.3255 | 0.37376 | 0.39938 | 0.37117 | 0.36267 | 0.37067 | 0.36 | 0.35638 | 0.38267 | 0.36577 | 0.35333 | 0.34445 | 0.37954 | 0.3768
| Median |  | 0.3245 | 0.37324 | 0.39946 | 0.371 | 0.364 | 0.372 | 0.36 | 0.35701 | 0.3825 | 0.3666 | 0.355 | 0.34259 | 0.3791 | 0.37
| Std Dev. |  | 0.0038859 | 0.0028231 | 0.0052139 | 0.0038166 | 0.0060222 | 0.003266 | 0.0043818 | 0.0023481 | 0.0029412 | 0.003563 | 0.002582 | 0.0039036 | 0.024687 | 0.020741
| Rel.Std.Dev. |  | 1.194% | 0.7553% | 1.306% | 1.028% | 1.661% | 0.8811% | 1.217% | 0.6589% | 0.7686% | 0.9741% | 0.7308% | 1.133% | 6.504% | 5.504%
| PDM3 |  | -11.49% | 1.628% | 8.594% | 0.9235% | -1.388% | 0.7875% | -2.113% | -3.099% | 4.05% | -0.5448% | -3.926% | -6.342% | 3.199% | 2.455%
| Z-Score (Absolute) |  | 2.86 | 0.38 | 2.1 | 0.21 | 0.36 | 0.16 | 0.54 | 0.79 | 0.98 | 0.16 | 0.99 | 1.59 | 0.77 | 0.58
| Analytical results for Mo in OREAS 196 (Indicative Value 1.12 ppm) | [16 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 11 | 12 | 14 | 15 | 17 | 20 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 1 | 1.0211 | 10 | 1 | 3 | < 2 | < 1 | 1.3333 | 2 | [6 empty] | 1
| 1 | 2 | 1 | 1.0364 | 7 | 1 | 3 | < 2 | < 1 | 1 | < 2 | [6 empty] | 4
| 1 | 3 | < 1 | 1.1393 | 9 |  | 3 | < 2 | < 1 | 1 | 2 | [6 empty] | 16
| 1 | 4 | 1 | 1.1816 | 8 |  | 3 | < 2 | < 1 | 1.5 | < 2 | [6 empty] | 1.1237
| 1 | 5 | 1 | < 1 | 8 |  | 3 | < 2 | < 1 | 2 | < 2 | [6 empty] | 21
| 1 | 6 | 1 | < 1 | 10 |  | 3 | < 2 | < 1 | 1 | 2
| Mean |  | 1 | 1.0946 | 8.6667 | 1 | 3 |  |  | 1.3056 | 2
| Median |  | 1 | 1.0879 | 8.5 | 1 | 3 |  |  | 1.1667 | 2
| Std Dev. |  | 0 | 0.078253 | 1.2111 | 0 | 0 |  |  | 0.40023 | 0
| Rel.Std.Dev. |  | 0% | 7.149% | 13.97% | 0% | 0% |  |  | 30.66% | 0%
| PDM3 |  | -11.01% | -2.592% | 671.2% | -11.01% | 167% |  |  | 16.18% | 77.98%
| Z-Score (Absolute) |  | 0.18 | 0.23 | 16.92 | 0 | 4.41 | 0 | 1.1 | 0.67 | 1.1
| Analytical results for Na in OREAS 196 (Indicative Value 0.139 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 14 | 15 | [13 empty] | wt.%
|  |  | BF*OES | PF*OES | [13 empty] | 3
3
| 1 | 1 | 0.09 | 0.17 | [13 empty] | 1
| 1 | 2 | 0.1 | 0.19 | [13 empty] | 16
| 1 | 3 | 0.1 | 0.17 | [13 empty] | 16
| 1 | 4 | 0.09 | 0.2 | [13 empty] | 0.13917
| 1 | 5 | 0.09 | 0.2 | [13 empty] | 22
| 1 | 6 | 0.09 | 0.18
| Mean |  | 0.093333 | 0.185
| Median |  | 0.09 | 0.185
| Std Dev. |  | 0.005164 | 0.013784
| Rel.Std.Dev. |  | 5.533% | 7.451%
| PDM3 |  | -32.93% | 32.93%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 196 (Indicative Value < 2 ppm) | [16 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 12 | 14 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 0.42105 | < 2 | < 2.4 | < 1 | < 5 | < 0.5 | 6.5 | < 10 | [7 empty] | 1
| 1 | 2 | 0.4936 | < 2 | 3 | < 1 | < 5 | < 0.5 | 14.75 | < 10 | [7 empty] | 6
| 1 | 3 | 0.41904 | < 2 | < 2.4 | < 1 | < 5 | < 0.5 | 7 | < 10 | [7 empty] | 16
| 1 | 4 | 0.53337 | < 2 | < 2.4 | < 1 | < 5 | < 0.5 | 12.6 | < 10 | [7 empty] | < 2
| 1 | 5 | 0.50643 | < 2 | < 2.4 | < 1 | < 5 | < 0.5 | 17 | < 10 | [7 empty] | 23
| 1 | 6 | 0.39714 | < 2 | < 2.4 | < 1 | < 5 | < 0.5 | 16.5 | < 10
| Mean |  | 0.46177 |  | 3 |  |  |  | 12.3917
| Median |  | 0.45732 |  | 3 |  |  |  | 13.675
| Std Dev. |  | 0.056206 | [5 empty] | 4.6369
| Rel.Std.Dev. |  | 12.17% | [5 empty] | 37.42%
PDM3
| Z-Score (Absolute) |  | 0.59 | 0.19 | 0.19 | 0.57 | 0.94 | 0.75 | 8.4 | 2.83
| Analytical results for Nd in OREAS 196 (Certified Value 9.25 ppm) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 05 | 12 | 14 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
3
| 1 | 1 | 10 | 9 | 8.6 | 9.2 | 9 | 10.62 | 8.5 | 11.2 | [7 empty] | 1
| 1 | 2 | 9 | 9.1 | 9.4 | 9.1 | 9.5 | 10.73 | 9.68 | 9.9 | [7 empty] | #N/A
| 1 | 3 | 10 | 9.2 | 8.8 | 8.9 | 9.5 | 10.93 | 10 | 10.5 | [7 empty] | 16
| 1 | 4 | 9 | 9.3 | 10.4 | 9.1 | 9.5 | 10.95 | 9.24 | 10.6 | [7 empty] | 9.2473
| 1 | 5 | 10 | 9.4 | 9.1 | 9.2 | 9 | 10.8 | 8.94 | 10.7 | [7 empty] | 51
| 1 | 6 | 9 | 9.3 | 9.3 | 9.5 | 9 | 10.94 | 9.66 | 11.5
| Mean |  | 9.5 | 9.2167 | 9.2667 | 9.1667 | 9.25 | 10.8283 | 9.3367 | 10.7333
| Median |  | 9.5 | 9.25 | 9.2 | 9.15 | 9.25 | 10.865 | 9.45 | 10.65
| Std Dev. |  | 0.54772 | 0.1472 | 0.6314 | 0.19664 | 0.27386 | 0.13527 | 0.55302 | 0.56095
| Rel.Std.Dev. |  | 5.766% | 1.597% | 6.814% | 2.145% | 2.961% | 1.249% | 5.923% | 5.226%
| PDM3 |  | 2.732% | -0.3316% | 0.2091% | -0.8723% | 0.02884% | 17.1% | 0.966% | 16.07%
| Z-Score (Absolute) |  | NA | 0.48 | 0 | 0.96 | 0.16 | 15.04 | 0.67 | 14.13
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 196 (Certified Value 2.02 wt.%) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 | wt.%
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | 1.813 | 1.94 | 1.9602 | 2.2515 | 2.05 | 1.985 | 2.06 | 1.885 | 2.2563 | 2.0697 | > 1 | 2.15 | 2.07 | 2.196 | 2.0973 | 1
| 1 | 2 | 1.864 | 1.97 | 1.9265 | 2.1786 | 2.04 | 2 | 2.01 | 1.88 | 2.2438 | 2.0507 | > 1 | 2.13 | 2.07 | 1.97 | 2.0367 | 22
| 1 | 3 | 1.912 | 1.97 | 1.8764 | 2.2201 | 2.05 | 2.01 | 2.05 | 1.895 | 2.2597 | 2.016 | > 1 | 2.14 | 2.07 | 2.174 | 1.9692 | 16
| 1 | 4 | 1.7924 | 1.94 | 1.9505 | 2.2367 | 2.05 | 1.96 | 2.07 | 1.89 | 2.2461 | 2.0557 | > 1 | 2.16 | 2.06 | 1.914 | 2.0047 | 2.0197
| 1 | 5 | 1.8323 | 2 | 1.9774 | 2.2769 | 2.03 | 2.02 | 2.05 | 1.905 | 2.2427 | 2.037 | > 1 | 2.13 | 2.06 | 1.918 | 2.0241 | 52
| 1 | 6 | 1.8243 | 1.96 | 1.9276 | 2.1923 | 2.03 | 2.06 | 2.07 | 1.88 | 2.2496 | 2.0386 | > 1 | 2.16 | 2.08 | 1.964 | 2.0007
| Mean |  | 1.8397 | 1.9633 | 1.9364 | 2.226 | 2.0417 | 2.0058 | 2.0517 | 1.8892 | 2.2497 | 2.0446 |  | 2.145 | 2.0683 | 2.0227 | 2.0221
| Median |  | 1.8283 | 1.965 | 1.9391 | 2.2284 | 2.045 | 2.005 | 2.055 | 1.8875 | 2.2478 | 2.0446 |  | 2.145 | 2.07 | 1.967 | 2.0144
| Std Dev. |  | 0.042567 | 0.022509 | 0.035258 | 0.036796 | 0.0098319 | 0.033826 | 0.022286 | 0.009704 | 0.0069395 | 0.018451 |  | 0.013784 | 0.0075277 | 0.128 | 0.043431
| Rel.Std.Dev. |  | 2.314% | 1.146% | 1.821% | 1.653% | 0.4816% | 1.686% | 1.086% | 0.5137% | 0.3085% | 0.9024% |  | 0.6426% | 0.364% | 6.328% | 2.148%
| PDM3 |  | -8.915% | -2.792% | -4.124% | 10.21% | 1.086% | -0.6882% | 1.581% | -6.465% | 11.39% | 1.232% |  | 6.202% | 2.406% | 0.1452% | 0.118%
| Z-Score (Absolute) |  | 2.47 | 0.88 | 1.23 | 2.49 | 0.12 | 0.34 | 0.25 | 1.84 | 2.79 | 0.16 | NA | 1.45 | 0.46 | 0.12 | 0.13
| Analytical results for P in OREAS 196 (Indicative Value 0.007 wt.%) | [16 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 12 | 14 | 15 | 17 | 20 | [8 empty] | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | [8 empty] | 3
3
| 1 | 1 | < 0.01 | 0.0027161 | < 0.01 | 0.015 | < 0.01 | 0.0023564 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 0.01 | 0.0028155 | < 0.01 | 0.02 | < 0.01 | 0.0013091 | 0.01 | [8 empty] | 18
| 1 | 3 | < 0.01 | 0.002952 | < 0.01 | 0.015 | < 0.01 | 0.0014546 | 0.01 | [8 empty] | 16
| 1 | 4 | < 0.01 | 0.0029896 | < 0.01 | 0.015 | < 0.01 | 0.0022692 | < 0.01 | [8 empty] | 0.0070002
| 1 | 5 | < 0.01 | 0.002666 | < 0.01 | 0.015 | < 0.01 | 0.0012219 | 0.01 | [8 empty] | 24
| 1 | 6 | < 0.01 | 0.002854 | < 0.01 | 0.01 | < 0.01 | 0.0024001 | 0.01
| Mean |  |  | 0.0028322 |  | 0.015 |  | 0.0018352 | 0.01
| Median |  |  | 0.0028347 |  | 0.015 |  | 0.0018619 | 0.01
| Std Dev. |  |  | 0.00012726 |  | 0.0031623 |  | 0.00056159 | 0
| Rel.Std.Dev. |  |  | 4.493% |  | 21.08% |  | 30.6% | 0%
| PDM3 |  |  | -59.54% |  | 114.3% |  | -73.78% | 42.85%
| Z-Score (Absolute) |  | 0 | 0.67 | 0 | 3.11 | 0 | 0.98 | 1.04
| Analytical results for Pb in OREAS 196 (Indicative Value 30.7 ppm) | [16 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 08 | 09 | 12 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | [5 empty] | 1
1
| 1 | 1 | < 20 | < 20 | 16 | 100 | < 100 | < 100 | < 5 | 27.8 |  | < 20 | [5 empty] | 1
| 1 | 2 | < 20 | < 20 | 17.2 | 100 | < 100 | < 100 | < 5 | 24.4 |  | < 20 | [5 empty] | 7
| 1 | 3 | < 20 | < 20 | 9 | < 100 | < 100 | < 100 | < 5 | 26 |  | < 20 | [5 empty] | 16
| 1 | 4 | < 20 | < 20 | 11.3 | 100 | < 100 | < 100 | < 5 | 25.4 |  | < 20 | [5 empty] | 30.7
| 1 | 5 | < 20 | < 20 | 14.4 | 100 | < 100 | < 100 | < 5 | 21.9 |  | < 20 | [5 empty] | 25
| 1 | 6 | < 20 | < 20 | 8.6 | < 100 | < 100 | < 100 | < 5 | 22.8 | 2 | < 20
| Mean |  |  |  | 12.75 | 100 |  |  |  | 24.7167 | 2
| Median |  |  |  | 12.85 | 100 |  |  |  | 24.9 | 2
| Std Dev. |  |  |  | 3.6462 | 0 |  |  |  | 2.16
| Rel.Std.Dev. |  |  |  | 28.6% | 0% |  |  |  | 8.739%
| PDM3 |  |  |  | -58.47% | 225.7% |  |  |  | -19.49% | -93.49%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 0.1 | 5.32 | 2.85 | 2.85 | 0.66 | 0.99 | 0.69 | 0.1
| Analytical results for Pr in OREAS 196 (Certified Value 2.39 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
3
| 1 | 1 | 3 | 2.17 | 2.5142 | 2.4 | 0.67246 | 2.41 | 3 | 2.5 | 2.33 | 2.84 | [5 empty] | 1
| 1 | 2 | 2 | 2.27 | 2.4546 | 2.6 | 0.60133 | 2.35 | 3 | 2.4 | 2.538 | 2.5 | [5 empty] | #N/A
| 1 | 3 | 3 | 2.25 | 2.3469 | 2.3 | 0.48642 | 2.35 | 3 | 2.5 | 2.826 | 2.62 | [5 empty] | 16
| 1 | 4 | 2 | 2.26 | 2.2082 | 2.5 | 0.45461 | 2.38 | 3 | 2.5 | 2.416 | 2.6 | [5 empty] | 2.387
| 1 | 5 | 3 | 2.2 | 2.3382 | 2.4 | 0.53043 | 2.4 | 3 | 2.4 | 2.386 | 2.75 | [5 empty] | 53
| 1 | 6 | 3 | 2.22 | 2.3045 | 2.5 | 0.68496 | 2.5 | 3 | 2.4 | 2.502 | 2.96
| Mean |  | 2.6667 | 2.2283 | 2.3611 | 2.45 | 0.5717 | 2.3983 | 3 | 2.45 | 2.4997 | 2.7117
| Median |  | 3 | 2.235 | 2.3425 | 2.45 | 0.56588 | 2.39 | 3 | 2.45 | 2.459 | 2.685
| Std Dev. |  | 0.5164 | 0.038687 | 0.10908 | 0.10488 | 0.096484 | 0.055648 | 0 | 0.054772 | 0.17702 | 0.17046
| Rel.Std.Dev. |  | 19.36% | 1.736% | 4.62% | 4.281% | 16.88% | 2.32% | 0% | 2.236% | 7.082% | 6.286%
| PDM3 |  | 11.71% | -6.648% | -1.086% | 2.638% | -76.05% | 0.4736% | 25.68% | 2.638% | 4.719% | 13.6%
| Z-Score (Absolute) |  | NA | 1.91 | 0.61 | 0.25 | 18.03 | 0.25 | NA | 0.25 | 0.73 | 2.8
Results from laboratories 1 and 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 196 (Certified Value 1.69 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 17 | 20 | [5 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 1.8 | 1.8623 | 0.89059 | 2.1 | < 0.5 | 1.9 | 1.8 | 2 |  | 1.3 | [5 empty] | 1
| 1 | 2 | 1.9 | 1.6451 | 0.94844 | 2.6 | < 0.5 | 1.5 | 1.5 | 2 | 1.75 | 1.6 | [5 empty] | 8
| 1 | 3 | 1.6 | 1.563 | 0.89559 | 2.2 | < 0.5 | 1.8 | 1.6 | 2 | 1.75 | 1.3 | [5 empty] | 16
| 1 | 4 | 1.9 | 1.9292 | 0.6898 | 2.1 | < 0.5 | 1.6 | 2 | 2 | 1.6 | 1.7 | [5 empty] | 1.6908
| 1 | 5 | 1.8 | 1.6509 | 0.66732 | 2.3 | < 0.5 | 1.4 | 2.2 | 2 | 2 | 0.9 | [5 empty] | 54
| 1 | 6 | 1.6 | 1.7184 | 0.8906 | 1.9 | < 0.5 | 1.5 | 1.6 | 2 | 1.75 | 1.5
| Mean |  | 1.7667 | 1.7282 | 0.83039 | 2.2 |  | 1.6167 | 1.7833 | 2 | 1.77 | 1.3833
| Median |  | 1.8 | 1.6847 | 0.8906 | 2.15 |  | 1.55 | 1.7 | 2 | 1.75 | 1.4
| Std Dev. |  | 0.13663 | 0.14045 | 0.11983 | 0.23664 |  | 0.19408 | 0.27142 | 0 | 0.14405 | 0.28577
| Rel.Std.Dev. |  | 7.734% | 8.127% | 14.43% | 10.76% |  | 12% | 15.22% | 0% | 8.138% | 20.66%
| PDM3 |  | 4.487% | 2.209% | -50.89% | 30.12% |  | -4.385% | 5.473% | 18.29% | 4.684% | -18.18%
| Z-Score (Absolute) |  | 0.23 | 0 | 5.43 | 2.85 | 8.94 | 0.67 | 0.33 | NA | 0.25 | 2.09
Results from laboratory 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Re in OREAS 196 (Certified Value < 0.1 ppm) | [16 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 02 | 04 | 14 | 20 | [11 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*MS | [11 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | [11 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | [11 empty] | #N/A
| 1 | 3 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | [11 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | [11 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | [11 empty] | 55
| 1 | 6 | < 0.1 | < 0.1 | < 0.1 | < 0.05
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for S in OREAS 196 (Certified Value 0.178 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 15 | 17 | 20 |  |  |  | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  | 3
3
| 1 | 1 | 0.1673 | 0.19 | 0.19294 | 0.18 | 0.19 | 0.21 | 0.17 | 0.023184 | 0.19 | 0.15 | 0.154 | 0.16 |  |  |  | 1
| 1 | 2 | 0.1661 | 0.19 | 0.18775 | 0.19 | 0.2 | 0.2 | 0.16 | 0.023296 | 0.19 | 0.14 | 0.172 | 0.15 |  |  |  | #N/A
| 1 | 3 | 0.1724 | 0.19 | 0.19817 | 0.19 | 0.2 | 0.2 | 0.18 | 0.023856 | 0.19 | 0.15 | 0.152 | 0.16 |  |  |  | 16
| 1 | 4 | 0.1613 | 0.19 | 0.19641 | 0.18 | 0.2 | 0.21 | 0.19 | 0.023632 | 0.18 | 0.14 | 0.184 | 0.15 |  |  |  | 0.17842
| 1 | 5 | 0.1692 | 0.2 | 0.18888 | 0.18 | 0.23 | 0.21 | 0.19 | 0.023408 | 0.18 | 0.15 | 0.148 | 0.15 |  |  |  | 56
| 1 | 6 | 0.1604 | 0.19 | 0.18314 | 0.19 | 0.22 | 0.21 | 0.16 | 0.023408 | 0.18 | 0.14 | 0.132 | 0.15
| Mean |  | 0.16612 | 0.19167 | 0.19121 | 0.185 | 0.20667 | 0.20667 | 0.175 | 0.023464 | 0.185 | 0.145 | 0.157 | 0.15333
| Median |  | 0.1667 | 0.19 | 0.19091 | 0.185 | 0.2 | 0.21 | 0.175 | 0.023408 | 0.185 | 0.145 | 0.153 | 0.15
| Std Dev. |  | 0.0046102 | 0.0040825 | 0.0056728 | 0.0054772 | 0.015055 | 0.005164 | 0.013784 | 0.00024281 | 0.0054772 | 0.0054772 | 0.018407 | 0.005164
| Rel.Std.Dev. |  | 2.775% | 2.13% | 2.967% | 2.961% | 7.285% | 2.499% | 7.877% | 1.035% | 2.961% | 3.777% | 11.72% | 3.368%
| PDM3 |  | -6.898% | 7.422% | 7.168% | 3.686% | 15.83% | 15.83% | -1.919% | -86.85% | 3.686% | -18.73% | -12.01% | -14.06%
| Z-Score (Absolute) |  | 0.51 | 0.43 | 0.41 | 0.18 | 0.98 | 0.98 | 0.18 | 5.72 | 0.18 | 1.28 | 0.84 | 0.98
| Analytical results for Sb in OREAS 196 (Indicative Value 0.65 ppm) | [16 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | < 2 | 48.91 |  | < 0.1 | < 1 | 52.9 |  | < 0.5 | [5 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | < 2 | 46.4 | 1.1 | < 0.1 | 2 | 53.2 |  | < 0.5 | [5 empty] | 9
| 1 | 3 | < 0.5 | < 0.5 | < 2 | 47.99 | 2.6 | < 0.1 | < 1 | 53.2 |  | 0.5 | [5 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | < 2 | 46.72 | 0.9 | < 0.1 | 1 | 53.1 | 0.4 | < 0.5 | [5 empty] | 0.64583
| 1 | 5 | < 0.5 | < 0.5 | < 2 | 45.72 | 0.9 | < 0.1 | < 1 | 54.9 | 0.2 | < 0.5 | [5 empty] | 20
| 1 | 6 | < 0.5 | < 0.5 | < 2 | 48.02 | 0.9 | < 0.1 | 2 | 52.2 |  | < 0.5
| Mean |  |  |  |  | 47.2933 | 1.28 |  | 1.6667 | 53.25 | 0.3 | 0.5
| Median |  |  |  |  | 47.355 | 0.9 |  | 2 | 53.15 | 0.3 | 0.5
| Std Dev. |  |  |  |  | 1.2024 | 0.74297 |  | 0.57735 | 0.89163 | 0.14142
| Rel.Std.Dev. |  |  |  |  | 2.542% | 58.04% |  | 34.64% | 1.674% | 47.14%
| PDM3 |  |  |  |  | 7223% | 98.19% |  | 158.1% | 8145% | -53.55% | -22.58%
| Z-Score (Absolute) |  | 0.65 | 0.65 | 0.57 | 75.48 | 1.02 | 0.97 | 0.7 | 85.12 | 0.57 | 0.58
| Analytical results for Sc in OREAS 196 (Certified Value 22 ppm) | [16 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 1
2
| 1 | 1 | 21 | 20 | 21.1639 | < 20 | 15.55 | 22 | 24 | 25 | 21.3 | 19.6 | 20 |  |  |  |  | 1
| 1 | 2 | 22 | 20 | 20.8544 | < 20 | 16.5 | 23 | 24 | 25 | 20.7 | 25.6 | 23 |  |  |  |  | 10
| 1 | 3 | 22 | 20 | 20.7598 | 20.3385 | 14.8125 | 22 | 24 | 25 | 21.8 | 21.2 | 21 |  |  |  |  | 16
| 1 | 4 | 22 | 20 | 20.8313 | 20.6895 | 15.1625 | 23 | 24 | 25 | 20.9 | 25.4 | 22 |  |  |  |  | 22.0178
| 1 | 5 | 22 | 20 | 20.8649 | 22.4087 | 15.8 | 23 | 24 | 25 | 22.1 | 23.6 | 20 |  |  |  |  | 57
| 1 | 6 | 21 | 20 | 20.8463 | 22.2513 | 17.7875 | 23 | 24 | 20 | 20.6 | 23.2 | 21
| Mean |  | 21.6667 | 20 | 20.8868 | 21.422 | 15.9354 | 22.6667 | 24 | 24.1667 | 21.2333 | 23.1 | 21.1667
| Median |  | 22 | 20 | 20.8504 | 21.4704 | 15.675 | 23 | 24 | 25 | 21.1 | 23.4 | 21
| Std Dev. |  | 0.5164 | 0 | 0.14083 | 1.0602 | 1.0746 | 0.5164 | 0 | 2.0412 | 0.6121 | 2.352 | 1.169
| Rel.Std.Dev. |  | 2.383% | 0% | 0.6742% | 4.949% | 6.744% | 2.278% | 0% | 8.447% | 2.883% | 10.18% | 5.523%
| PDM3 |  | -1.595% | -9.164% | -5.137% | -2.706% | -27.62% | 2.947% | 9.003% | 9.76% | -3.563% | 4.915% | -3.865%
| Z-Score (Absolute) |  | 0.1 | NA | 0.26 | 1.8 | 2.59 | 0.57 | 1.2 | 1.28 | 0.1 | 0.78 | 0.13
Results from laboratory 2 were removed due to their 10 ppm reading resolution.
| Analytical results for Se in OREAS 196 (Certified Value < 20 ppm) | [16 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 04 | 05 | 14 | 17 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | BF*MS | PF*OES/MS | PF*MS | [9 empty] | 1
1
| 1 | 1 | < 20 | < 20 | 9 | < 10 | 3.5 | < 50 | [9 empty] | 1
| 1 | 2 | < 20 | < 20 | < 8 | < 10 | 5.2 | < 50 | [9 empty] | 1
| 1 | 3 | < 20 | < 20 | < 8 | < 10 | 2.4 | < 50 | [9 empty] | 16
| 1 | 4 | < 20 | < 20 | < 8 | < 10 | 3.5 | < 50 | [9 empty] | < 20
| 1 | 5 | < 20 | < 20 | < 8 | < 10 | 4 | < 50 | [9 empty] | 58
| 1 | 6 | < 20 | < 20 | < 8 | < 10 | 3 | < 50
| Mean |  |  |  | 9 |  | 3.6
| Median |  |  |  | 9 |  | 3.5
| Std Dev. | [5 empty] | 0.95289
| Rel.Std.Dev. | [5 empty] | 26.47%
PDM3
| Z-Score (Absolute) |  | 0.65 | 0.65 | 0.7 | 0.65 | 1.02 | 4.57
| Analytical results for Si in OREAS 196 (Certified Value 19.22 wt.%) | [16 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 17 | 20 |  |  | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES |  |  | 2
3
| 1 | 1 | 18.5 | 20.4443 | 20.1123 | 20 | 18.98 | 20.1 | 18.61 | 18.5897 | 19.8 | 19 | 19.3 | 21.2696 | 18.7 |  |  | 1
| 1 | 2 | 18.8 | 18.5226 | 19.744 | 19.3 | 19.4 | 19.4 | 18.61 | 18.8744 | 19.8 | 18.9 | 18.9 | 19.9046 | 19.3 |  |  | #N/A
| 1 | 3 | 18.8 | 20.0055 | 19.8931 | 19.3 | 19.4 | 19.68 | 18.61 | 18.6079 | 19.2 | 18.6 | 19 | 18.8669 | 18.6 |  |  | 16
| 1 | 4 | 19.2 | 20.4481 | 19.9875 | 19.2 | 19.21 | 20.01 | 18.79 | 18.5687 | 19.6 | 19 | 18.8 | 19.6709 | 18.5 |  |  | 19.2205
| 1 | 5 | 18.9 | 20.2701 | 20.1332 | 19.3 | 18.98 | 19.82 | 18.89 | 18.6501 | 19.6 | 19 | 18.7 | 18.6425 | 19.2 |  |  | 59
| 1 | 6 | 18.6 | 20.3963 | 19.744 | 19.3 | 18.98 | 20.01 | 18 | 18.6598 | 19.6 | 18.8 | 18.8 | 18.1563 | 19.2
| Mean |  | 18.8 | 20.0145 | 19.9357 | 19.4 | 19.1583 | 19.8367 | 18.585 | 18.6584 | 19.6 | 18.8833 | 18.9167 | 19.4185 | 18.9167
| Median |  | 18.8 | 20.3332 | 19.9403 | 19.3 | 19.095 | 19.915 | 18.61 | 18.629 | 19.6 | 18.95 | 18.85 | 19.2689 | 18.95
| Std Dev. |  | 0.24495 | 0.74961 | 0.17215 | 0.29665 | 0.20731 | 0.26265 | 0.30956 | 0.11139 | 0.21909 | 0.16021 | 0.2137 | 1.1159 | 0.35449
| Rel.Std.Dev. |  | 1.303% | 3.745% | 0.8635% | 1.529% | 1.082% | 1.324% | 1.666% | 0.597% | 1.118% | 0.8484% | 1.13% | 5.747% | 1.874%
| PDM3 |  | -2.188% | 4.131% | 3.721% | 0.9339% | -0.3234% | 3.206% | -3.306% | -2.924% | 1.974% | -1.754% | -1.581% | 1.03% | -1.581%
| Z-Score (Absolute) |  | 0.67 | 1.61 | 1.46 | 0.46 | 0 | 1.28 | 1.09 | 0.94 | 0.83 | 0.52 | 0.45 | 0.49 | 0.45
| Analytical results for Sm in OREAS 196 (Certified Value 1.63 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [6 empty] | ppm
|  |  | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | < 2 | 1.5545 | 1.5 | 1.6694 | 1.7 | 1.5 | < 1 | 1.86 | 2 | [6 empty] | 1
| 1 | 2 | < 2 | 1.6405 | 1.6 | 1.4909 | 1.7 | 1.5 | < 1 | 1.66 | 1.6 | [6 empty] | #N/A
| 1 | 3 | < 2 | 1.4932 | 1.3 | 1.5487 | 1.5 | 1.5 | < 1 | 1.9 | 1.8 | [6 empty] | 16
| 1 | 4 | < 2 | 1.5263 | 1.6 | 1.6956 | 1.7 | 1.5 | < 1 | 1.58 | 1.8 | [6 empty] | 1.6303
| 1 | 5 | < 2 | 1.4662 | 1.7 | 1.6431 | 1.7 | 1.5 | < 1 | 1.62 | 2 | [6 empty] | 60
| 1 | 6 | < 2 | 1.552 | 1.5 | 1.6508 | 1.7 | 1.5 | < 1 | 1.62 | 1.9
| Mean |  |  | 1.5388 | 1.5333 | 1.6164 | 1.6667 | 1.5 |  | 1.7067 | 1.85
| Median |  |  | 1.5391 | 1.55 | 1.6469 | 1.7 | 1.5 |  | 1.64 | 1.85
| Std Dev. |  |  | 0.060425 | 0.13663 | 0.079135 | 0.08165 | 0 |  | 0.13721 | 0.15166
| Rel.Std.Dev. |  |  | 3.927% | 8.91% | 4.896% | 4.899% | 0% |  | 8.04% | 8.198%
| PDM3 |  |  | -5.612% | -5.946% | -0.8495% | 2.233% | -7.99% |  | 4.686% | 13.48%
| Z-Score (Absolute) |  | 2.84 | 0 | 0.03 | 0.41 | 0.67 | 0.2 | 5.48 | 0.89 | 1.64
| Analytical results for Sn in OREAS 196 (Indicative Value 0.98 ppm) | [16 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 12 | 14 | 15 | 17 | 20 | [6 empty] | ppm
|  |  | PF*MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 6 | < 2 | < 100 | 1 | < 1 | < 10 | < 5 | 0.75 | < 10 | [6 empty] | 1
| 1 | 2 | < 5 | < 2 | < 100 | 1.2 | < 1 | < 10 | < 5 | 0.95 | < 10 | [6 empty] | 15
| 1 | 3 | 5 | < 2 | < 100 | 0.9 | < 1 | < 10 | < 5 | 0.6 | < 10 | [6 empty] | 16
| 1 | 4 | 6 | < 2 | < 100 | 1.1 | < 1 | < 10 | < 5 | 0.84 | < 10 | [6 empty] | 0.9775
| 1 | 5 | 6 | < 2 | < 100 | 1.3 | < 1 | < 10 | < 5 | 1.44 | < 10 | [6 empty] | 21
| 1 | 6 | 8 | < 2 | < 100 | 1.1 | < 1 | < 10 | < 5 | 0.55 | < 10
| Mean |  | 6.2 |  |  | 1.1 |  |  |  | 0.855
| Median |  | 6 |  |  | 1.1 |  |  |  | 0.795
| Std Dev. |  | 1.0954 |  |  | 0.14142 |  |  |  | 0.32266
| Rel.Std.Dev. |  | 17.67% |  |  | 12.86% |  |  |  | 37.74%
| PDM3 |  | 534.3% |  |  | 12.53% |  |  |  | -12.53%
| Z-Score (Absolute) |  | 1.04 | 0.51 | 16.01 | 0.47 | 0.67 | 0.84 | 0 | 0.55 | 0.84
| Analytical results for Sr in OREAS 196 (Certified Value 19.7 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | BF*MS | PF*OES | PF*OES/MS | PF*MS |  |  |  |  | 1
1
| 1 | 1 | < 20 | 21.3159 | < 20 | 30 | 19.13 | 23 | 20 | 20 | 17.1 | 21.6 | 19.2 |  |  |  |  | 1
| 1 | 2 | < 20 | 20.8039 | < 20 | 29 | 17.99 | 23 | 18 | 20 | 16.6 | 20 | 18.2 |  |  |  |  | #N/A
| 1 | 3 | < 20 | 22.4299 | < 20 | 27 | 18.75 | 23 | 19 | 20 | 16.6 | 18.2 | 19.9 |  |  |  |  | 16
| 1 | 4 | < 20 | 19.3826 | < 20 | 29 | 19.52 | 23 | 19 | 20 | 16.6 | 21.4 | 17.5 |  |  |  |  | 19.6984
| 1 | 5 | < 20 | 23.0379 | < 20 | 28 | 20.31 | 23 | 18 | 20 | 17.1 | 18.6 | 19 |  |  |  |  | 61
| 1 | 6 | < 20 | 22.8636 | < 20 | 27 | 18.2 | 23 | 19 | 20 | 16.6 | 16.8 | 21.6
| Mean |  |  | 21.639 |  | 28.3333 | 18.9833 | 23 | 18.8333 | 20 | 16.7667 | 19.4333 | 19.2333
| Median |  |  | 21.8729 |  | 28.5 | 18.94 | 23 | 19 | 20 | 16.6 | 19.3 | 19.1
| Std Dev. |  |  | 1.4125 |  | 1.2111 | 0.86306 | 0 | 0.75277 | 0 | 0.2582 | 1.8991 | 1.4264
| Rel.Std.Dev. |  |  | 6.527% |  | 4.274% | 4.546% | 0% | 3.997% | 0% | 1.54% | 9.773% | 7.416%
| PDM3 |  |  | 9.851% |  | 43.84% | -3.63% | 16.76% | -4.392% | 1.531% | -14.88% | -1.346% | -2.361%
| Z-Score (Absolute) |  | 2.52 | 0.7 | 2.52 | 2.55 | 0.03 | 1.08 | 0.08 | NA | 0.65 | 0.09 | 0.03
| Analytical results for Ta in OREAS 196 (Indicative Value 0.52 ppm) | [16 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | 2 | 0.9 | < 0.2 | < 0.1 | 0.5 | 0.1 | < 0.5 | < 2 | < 0.5 | 0.08 | < 0.5 |  |  |  |  | 1
| 1 | 2 | 0.3 | 0.8 | < 0.2 | < 0.1 | 1.4 |  | < 0.5 | < 2 | < 0.5 |  | < 0.5 |  |  |  |  | 3
| 1 | 3 | 0.4 | 0.9 | < 0.2 | < 0.1 | 0.6 | 0.2 | < 0.5 | < 2 | < 0.5 | 0.055 | < 0.5 |  |  |  |  | 16
| 1 | 4 | 2 | 0.8 | < 0.2 | < 0.1 | 0.6 | 0.1 | < 0.5 | < 2 | < 0.5 | 0.065 | < 0.5 |  |  |  |  | 0.51773
| 1 | 5 | 0.4 | 0.8 | < 0.2 | < 0.1 | 0.6 | 0.4 | < 0.5 | < 2 | < 0.5 | 0.06 | < 0.5 |  |  |  |  | 22
| 1 | 6 | 0.5 | 0.8 | < 0.2 | < 0.1 | 0.5 |  | < 0.5 | < 2 | < 0.5 | 0.05 | < 0.5
| Mean |  | 0.93333 | 0.83333 |  |  | 0.7 | 0.2 |  |  |  | 0.062
| Median |  | 0.45 | 0.8 |  |  | 0.6 | 0.15 |  |  |  | 0.06
| Std Dev. |  | 0.82865 | 0.05164 |  |  | 0.34641 | 0.14142 |  |  |  | 0.011511
| Rel.Std.Dev. |  | 88.78% | 6.197% |  |  | 49.49% | 70.71% |  |  |  | 18.57%
| PDM3 |  | 80.27% | 60.96% |  |  | 35.2% | -61.37% |  |  |  | -88.02%
| Z-Score (Absolute) |  | 2.45 | 2.09 | 0.54 | 0.72 | 1.61 | 0.18 | 0 | 2.69 | 0 | 0.67 | 0
| Analytical results for Tb in OREAS 196 (Certified Value 0.25 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 12 | 14 | 15 | 17 | 20 | [7 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 2
3
| 1 | 1 | 0.24 | 0.24182 | 0.2 | 0.24 | < 0.5 | < 0.5 | 0.306 | 0.29 | [7 empty] | 1
| 1 | 2 | 0.24 | 0.25851 | 0.2 | 0.23 | < 0.5 | < 0.5 | 0.252 | 0.28 | [7 empty] | #N/A
| 1 | 3 | 0.26 | 0.24698 | 0.3 | 0.23 | < 0.5 | < 0.5 | 0.222 | 0.28 | [7 empty] | 16
| 1 | 4 | 0.25 | 0.23221 | 0.3 | 0.22 | < 0.5 | < 0.5 | 0.236 | 0.28 | [7 empty] | 0.25437
| 1 | 5 | 0.26 | 0.27275 | 0.3 | 0.25 | < 0.5 | < 0.5 | 0.244 | 0.31 | [7 empty] | 62
| 1 | 6 | 0.25 | 0.23691 | 0.2 | 0.25 | < 0.5 | < 0.5 | 0.256 | 0.33
| Mean |  | 0.25 | 0.2482 | 0.25 | 0.23667 |  |  | 0.25267 | 0.295
| Median |  | 0.25 | 0.2444 | 0.25 | 0.235 |  |  | 0.248 | 0.285
| Std Dev. |  | 0.0089443 | 0.015063 | 0.054772 | 0.012111 |  |  | 0.028807 | 0.020736
| Rel.Std.Dev. |  | 3.578% | 6.069% | 21.91% | 5.117% |  |  | 11.4% | 7.029%
| PDM3 |  | -1.719% | -2.428% | -1.719% | -6.961% |  |  | -0.6707% | 15.97%
| Z-Score (Absolute) |  | 0 | 0.67 | NA | 4.99 | 0 | 0 | 1 | 16.83
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 196 (Certified Value < 1 ppm) | [16 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 12 | 14 | 17 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | < 1 | < 2 | < 6 | < 1 | < 5 | 0.66 | < 1 | [8 empty] | 1
| 1 | 2 | < 1 | < 2 | < 6 | < 1 | < 5 | 1.225 | < 1 | [8 empty] | 4
| 1 | 3 | < 1 | < 2 | < 6 | < 1 | < 5 | 1.28 | < 1 | [8 empty] | 16
| 1 | 4 | < 1 | < 2 | < 6 | < 1 | < 5 | 1.38 | < 1 | [8 empty] | < 1
| 1 | 5 | < 1 | < 2 | < 6 | < 1 | < 5 | 1.5 | < 1 | [8 empty] | 63
| 1 | 6 | < 1 | < 2 | < 6 | < 1 | < 5 | 0.625 | < 1
| Mean | [6 empty] | 1.1117
| Median | [6 empty] | 1.2525
| Std Dev. | [6 empty] | 0.37546
| Rel.Std.Dev. | [6 empty] | 33.77%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 2.7 | 0.67 | 2.02 | 0.15 | 0.67
| Analytical results for Th in OREAS 196 (Certified Value 0.19 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 11 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 0.2 | < 0.2 | 0.11909 | 0.2 | 0.2 | 0.2 | < 0.5 | < 0.5 | 0.1 | 0.2 | [5 empty] | 1
| 1 | 2 | 0.2 | < 0.2 | 0.11209 | 0.2 | 0.1 | 0.2 | < 0.5 | < 0.5 | 0.14 | 0.2 | [5 empty] | #N/A
| 1 | 3 | 0.2 | < 0.2 | < 0.1 | 0.2 | 0.2 | 0.2 | < 0.5 | < 0.5 | 0.16 | 0.2 | [5 empty] | 16
| 1 | 4 | 0.3 | < 0.2 | 0.10155 | 0.2 | 0.2 | 0.2 | < 0.5 | < 0.5 | 0.14 | 0.2 | [5 empty] | 0.18778
| 1 | 5 | 0.2 | < 0.2 | 0.11576 | 0.2 | 0.2 | 0.2 | < 0.5 | < 0.5 | 0.16 | 0.2 | [5 empty] | 64
| 1 | 6 | 0.2 | < 0.2 | < 0.1 | 0.2 | 0.2 | 0.2 | < 0.5 | < 0.5 | 0.16 | 0.2
| Mean |  | 0.21667 |  | 0.11212 | 0.2 | 0.18333 | 0.2 |  |  | 0.14333 | 0.2
| Median |  | 0.2 |  | 0.11392 | 0.2 | 0.2 | 0.2 |  |  | 0.15 | 0.2
| Std Dev. |  | 0.040825 |  | 0.007609 | 3.0405e-17 | 0.040825 | 3.0405e-17 |  |  | 0.023381 | 3.0405e-17
| Rel.Std.Dev. |  | 18.84% |  | 6.786% | 1.52e-14% | 22.27% | 1.52e-14% |  |  | 16.31% | 1.52e-14%
| PDM3 |  | 15.38% |  | -40.29% | 6.509% | -2.367% | 6.509% |  |  | -23.67% | 6.509%
| Z-Score (Absolute) |  | 0.34 | 2.02 | 2.2 | 0 | 0.34 | 0 | 1.01 | 1.01 | 1.15 | 0
| Analytical results for Ti in OREAS 196 (Certified Value 0.081 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  | wt.%
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.08 | 0.082825 | 0.083148 | 0.08 | 0.078 | 0.084 | 0.072 | 0.085296 | 0.08 | 0.08 | 0.08 | 0.07 | 0.098 | 0.08 |  | 1
| 1 | 2 | 0.08 | 0.081745 | 0.080112 | 0.08 | 0.072 | 0.078 | 0.102 | 0.085823 | 0.08 | 0.08 | 0.08 | 0.07 | 0.086 | 0.09 |  | #N/A
| 1 | 3 | 0.08 | 0.083115 | 0.085321 | 0.08 | 0.072 | 0.078 | 0.138 | 0.086209 | 0.08 | 0.08 | 0.08 | 0.07 | 0.068 | 0.08 |  | 16
| 1 | 4 | 0.08 | 0.086272 | 0.086483 | 0.08 | 0.078 | 0.078 | 0.15 | 0.085337 | 0.08 | 0.08 | 0.08 | 0.07 | 0.088 | 0.08 |  | 0.080561
| 1 | 5 | 0.08 | 0.084095 | 0.08572 | 0.08 | 0.072 | 0.078 | 0.072 | 0.085886 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | 0.08 |  | 65
| 1 | 6 | 0.08 | 0.08158 | 0.081609 | 0.08 | 0.072 | 0.078 | 0.102 | 0.085202 | 0.08 | 0.08 | 0.075 | 0.07 | 0.08 | 0.08
| Mean |  | 0.08 | 0.083272 | 0.083732 | 0.08 | 0.074 | 0.079 | 0.106 | 0.085625 | 0.08 | 0.08 | 0.079167 | 0.07 | 0.083333 | 0.081667
| Median |  | 0.08 | 0.08297 | 0.084235 | 0.08 | 0.072 | 0.078 | 0.102 | 0.08558 | 0.08 | 0.08 | 0.08 | 0.07 | 0.083 | 0.08
| Std Dev. |  | 0 | 0.0017377 | 0.0025296 | 0 | 0.0030984 | 0.0024495 | 0.03257 | 0.00040467 | 0 | 0 | 0.0020412 | 0 | 0.010013 | 0.0040825
| Rel.Std.Dev. |  | 0% | 2.087% | 3.021% | 0% | 4.187% | 3.101% | 30.73% | 0.4726% | 0% | 0% | 2.578% | 0% | 12.02% | 4.999%
| PDM3 |  | -0.6963% | 3.365% | 3.936% | -0.6963% | -8.144% | -1.938% | 31.58% | 6.287% | -0.6963% | -0.6963% | -1.731% | -13.11% | 3.441% | 1.372%
| Z-Score (Absolute) |  | 0 | 0.89 | 1.02 | 0 | 1.66 | 0.29 | 7.07 | 1.54 | 0 | 0 | 0.23 | 2.73 | 0.91 | 0.46
| Analytical results for Tl in OREAS 196 (Certified Value < 0.5 ppm) | [16 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 12 | 14 | 15 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 3 | < 1 | < 0.2 | [8 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 3 | < 1 | < 0.2 | [8 empty] | 5
| 1 | 3 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 3 | < 1 | < 0.2 | [8 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 3 | < 1 | < 0.2 | [8 empty] | < 0.5
| 1 | 5 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 3 | < 1 | < 0.2 | [8 empty] | 66
| 1 | 6 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 3 | < 1 | < 0.2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 0.9 | 0 | 5.62 | 1.12 | 0.67
| Analytical results for Tm in OREAS 196 (Certified Value 0.13 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | < 2 | 0.07 | 0.10945 | 0.1 | 0.1246 | 0.13 | < 1 | < 1 | 0.178 | 0.16 | [5 empty] | 1
| 1 | 2 | < 2 | 0.09 | 0.12276 | 0.1 | 0.1379 | 0.13 | < 1 | < 1 | 0.172 | 0.15 | [5 empty] | #N/A
| 1 | 3 | < 2 | 0.08 | 0.12257 | 0.1 | 0.1104 | 0.13 | < 1 | < 1 | 0.164 | 0.16 | [5 empty] | 16
| 1 | 4 | < 2 | 0.11 | 0.13529 | 0.1 | 0.1212 | 0.13 | < 1 | < 1 | 0.168 | 0.15 | [5 empty] | 0.13292
| 1 | 5 | < 2 | 0.11 | 0.12064 | 0.1 | 0.1062 | 0.13 | < 1 | < 1 | 0.17 | 0.16 | [5 empty] | 67
| 1 | 6 | < 2 | 0.08 | 0.12421 | 0.2 | 0.1298 | 0.14 | < 1 | < 1 | 0.178 | 0.18
| Mean |  |  | 0.09 | 0.12249 | 0.11667 | 0.12168 | 0.13167 |  |  | 0.17167 | 0.16
| Median |  |  | 0.085 | 0.12266 | 0.1 | 0.1229 | 0.13 |  |  | 0.171 | 0.16
| Std Dev. |  |  | 0.016733 | 0.0082506 | 0.040825 | 0.011871 | 0.0040825 |  |  | 0.0055737 | 0.010954
| Rel.Std.Dev. |  |  | 18.59% | 6.736% | 34.99% | 9.756% | 3.101% |  |  | 3.247% | 6.847%
| PDM3 |  |  | -32.29% | -7.848% | -12.23% | -8.452% | -0.9408% |  |  | 29.15% | 20.38%
| Z-Score (Absolute) |  | 14.78 | 1.23 | 0.66 | NA | 0.67 | 0.5 | 5.98 | 5.98 | 0.21 | 0
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 196 (Indicative Value 0.18 ppm) | [16 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 11 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | 0.3 | < 5 | 0.10134 | 0.2 | 0.1 | 0.12 | < 0.5 | < 0.5 | 0.3025 | < 2 | [5 empty] | 1
| 1 | 2 | 0.3 | < 5 | 0.11794 | 0.2 | 0.2 | 0.1 | < 0.5 | < 0.5 | 0.5775 | < 2 | [5 empty] | 17
| 1 | 3 | 0.3 | < 5 | 0.10191 | 0.2 | 0.2 | 0.11 | < 0.5 | < 0.5 | 0.39 | < 2 | [5 empty] | 16
| 1 | 4 | 0.3 | < 5 | 0.1254 | 0.2 | 0.2 | 0.11 | < 0.5 | < 0.5 | 0.6825 | < 2 | [5 empty] | 0.18386
| 1 | 5 | 0.3 | < 5 | < 0.1 | 0.2 | 0.4 | 0.1 | < 0.5 | < 0.5 | 0.8075 | < 2 | [5 empty] | 23
| 1 | 6 | 0.3 | < 5 | 0.11668 | 0.2 | 0.1 | 0.1 | < 0.5 | < 0.5 | 0.5 | < 2
| Mean |  | 0.3 |  | 0.11265 | 0.2 | 0.2 | 0.10667 |  |  | 0.54333
| Median |  | 0.3 |  | 0.11668 | 0.2 | 0.2 | 0.105 |  |  | 0.53875
| Std Dev. |  | 0 |  | 0.010607 | 3.0405e-17 | 0.10954 | 0.008165 |  |  | 0.18639
| Rel.Std.Dev. |  | 0% |  | 9.416% | 1.52e-14% | 54.77% | 7.655% |  |  | 34.31%
| PDM3 |  | 63.16% |  | -38.73% | 8.776% | 8.776% | -41.99% |  |  | 195.5%
| Z-Score (Absolute) |  | 0.35 | 15.7 | 1.03 | 0.35 | 0.35 | 1 | 0 | 0 | 2.05 | 5.23
| Analytical results for V in OREAS 196 (Certified Value 94 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | [5 empty] | 0
1
| 1 | 1 | 100 | 89.6331 | 96 | 84.5384 | 94 | 97 | 80 | 62.9 | 100 | 86 | [5 empty] | 1
| 1 | 2 | 100 | 95.6812 | 100 | 86.417 | 93 | 93 | 80 | 62 | 112.6 | 85 | [5 empty] | #N/A
| 1 | 3 | 100 | 96.6127 | 99 | 84.8825 | 93 | 93 | 80 | 64.4 | 106.4 | 83 | [5 empty] | 16
| 1 | 4 | 100 | 92.0327 | 97 | 84.1013 | 91 | 94 | 80 | 63.9 | 115 | 84 | [5 empty] | 93.5686
| 1 | 5 | 100 | 92.8283 | 99 | 85.5335 | 90 | 95 | 80 | 64.6 | 104.6 | 85 | [5 empty] | 68
| 1 | 6 | 100 | 90.1278 | 100 | 85.4934 | 90 | 96 | 80 | 64 | 105.4 | 85
| Mean |  | 100 | 92.8193 | 98.5 | 85.161 | 91.8333 | 94.6667 | 80 | 63.6333 | 107.3333 | 84.6667
| Median |  | 100 | 92.4305 | 99 | 85.188 | 92 | 94.5 | 80 | 63.95 | 105.9 | 85
| Std Dev. |  | 0 | 2.8494 | 1.6432 | 0.82643 | 1.7224 | 1.633 | 0 | 0.99331 | 5.5204 | 1.0328
| Rel.Std.Dev. |  | 0% | 3.07% | 1.668% | 0.9704% | 1.876% | 1.725% | 0% | 1.561% | 5.143% | 1.22%
| PDM3 |  | 6.873% | -0.8008% | 5.27% | -8.985% | -1.855% | 1.174% | -14.5% | -31.99% | 14.71% | -9.514%
| Z-Score (Absolute) |  | NA | 0.05 | 0.62 | 0.72 | 0.05 | 0.24 | NA | 2.9 | 1.52 | 0.77
Results from laboratories 2 and 14 were removed due to their 10 ppm reading resolution.
| Analytical results for W in OREAS 196 (Certified Value 2.21 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | 30000 | < 5 | 1.7385 | 3.1 | 5.1335 | 1 | 2 | < 3 | < 5 | 2 | 3 |  |  |  |  | 1
| 1 | 2 | 30000 | < 5 | 1.5735 | 2.6 | 5.2602 | 2 | 2 | < 3 | < 5 | 1.8 | < 2 |  |  |  |  | 6
| 1 | 3 | 30000 | < 5 | 1.5273 | 2.5 | 5.28 | 2 | 2 | 3 | < 5 | 2 | < 2 |  |  |  |  | 16
| 1 | 4 | 30000 | < 5 | 1.6985 | 1.8 | 5.2296 | 1 | 2 | < 3 | < 5 | 2 | < 2 |  |  |  |  | 2.2119
| 1 | 5 | 30000 | < 5 | 1.7864 | 2.6 | 5.2374 | 2 | 2 | < 3 | < 5 | 2 | < 2 |  |  |  |  | 69
| 1 | 6 | 30000 | < 5 | 1.6738 | 1.7 | 4.9622 | 1 | 2 | < 3 | < 5 | 1.8 | < 2
| Mean |  | 30000 |  | 1.6663 | 2.3833 | 5.1838 | 1.5 | 2 | 3 |  | 1.9333 | 3
| Median |  | 30000 |  | 1.6861 | 2.55 | 5.2335 | 1.5 | 2 | 3 |  | 2 | 3
| Std Dev. |  | 0 |  | 0.098654 | 0.53448 | 0.11975 | 0.54772 | 0 |  |  | 0.10328
| Rel.Std.Dev. |  | 0% |  | 5.92% | 22.43% | 2.31% | 36.51% | 0% |  |  | 5.342%
| PDM3 |  | 1356226% |  | -24.66% | 7.753% | 134.4% | -32.18% | -9.578% | 35.63% |  | -12.59% | 35.63%
| Z-Score (Absolute) |  | NA | 0.81 | 0.46 | 0.64 | 4.91 | 0.71 | 0.05 | 0.33 | 0.81 | 0.05 | 0.97
| Analytical results for Y in OREAS 196 (Certified Value 8.64 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 10 | 11 | 12 | 14 | 15 | 17 | 20 |  |  |  | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 9 | 8.4 | 8.5216 | 9.3614 | 9.1 | 8.16 | 8.6 | 8.4 | 9 | 9 | 8.18 | 8.6 |  |  |  | 1
| 1 | 2 | 9 | 8.5 | 8.272 | 8.6642 | 8.8 | 8.507 | 8.6 | 8.3 | 9 | 9.3 | 7.28 | 8 |  |  |  | 7
| 1 | 3 | 9 | 8.5 | 8.4375 | 9.153 | 9 | 8.359 | 8.7 | 8.2 | 9 | 8.9 | 5.54 | 8.6 |  |  |  | 16
| 1 | 4 | 9 | 8.5 | 7.8643 | 9.4849 | 8.6 | 8.308 | 8.9 | 8.5 | 9 | 9.1 | 7.175 | 8.2 |  |  |  | 8.6392
| 1 | 5 | 9 | 8.5 | 7.8252 | 9.7742 | 8.2 | 8.209 | 8.9 | 8.3 | 9 | 9.1 | 8.3 | 8.4 |  |  |  | 70
| 1 | 6 | 9 | 8.4 | 8.7889 | 8.905 | 9.3 | 8.22 | 8.8 | 8.6 | 9 | 8.7 | 7.12 | 9.2
| Mean |  | 9 | 8.4667 | 8.2849 | 9.2238 | 8.8333 | 8.2938 | 8.75 | 8.3833 | 9 | 9.0167 | 7.2658 | 8.5
| Median |  | 9 | 8.5 | 8.3547 | 9.2572 | 8.9 | 8.264 | 8.75 | 8.35 | 9 | 9.05 | 7.2275 | 8.5
| Std Dev. |  | 0 | 0.05164 | 0.37992 | 0.40244 | 0.39328 | 0.12684 | 0.13784 | 0.1472 | 0 | 0.20412 | 0.9913 | 0.41473
| Rel.Std.Dev. |  | 0% | 0.6099% | 4.586% | 4.363% | 4.452% | 1.529% | 1.575% | 1.756% | 0% | 2.264% | 13.64% | 4.879%
| PDM3 |  | 4.177% | -1.997% | -4.101% | 6.767% | 2.247% | -3.997% | 1.283% | -2.961% | 4.177% | 4.37% | -15.9% | -1.611%
| Z-Score (Absolute) |  | NA | 0.05 | 0.58 | 2.15 | 1.01 | 0.55 | 0.77 | 0.29 | NA | 1.55 | 3.53 | 0.05
Results from laboratories 1 and 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 196 (Certified Value 0.84 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 10 | 12 | 14 | 15 | 17 | 20 | [5 empty] | ppm
|  |  | PF*MS | PF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | < 2 | 0.7 | 0.89612 | 0.7 | 0.7158 | 0.8 | < 1 | 0.9 | 0.72 | 1 | [5 empty] | 1
| 1 | 2 | < 2 | 0.7 | 1.0027 | 0.7 | 0.7618 | 0.8 | < 1 | 0.9 | 0.92 | 0.9 | [5 empty] | #N/A
| 1 | 3 | < 2 | 0.7 | 0.85572 | 0.8 | 0.8226 | 0.8 | < 1 | 0.9 | 0.74 | 0.8 | [5 empty] | 16
| 1 | 4 | < 2 | 0.8 | 1.0005 | 0.8 | 0.797 | 0.9 | < 1 | 1 | 0.82 | 0.9 | [5 empty] | 0.8432
| 1 | 5 | < 2 | 0.8 | 0.89006 | 0.9 | 0.7848 | 0.8 | < 1 | 0.9 | 0.88 | 0.9 | [5 empty] | 71
| 1 | 6 | < 2 | 0.8 | 0.80248 | 1 | 0.7838 | 0.8 | < 1 | 0.9 | 0.88 | 1.1
| Mean |  |  | 0.75 | 0.90793 | 0.81667 | 0.77763 | 0.81667 |  | 0.91667 | 0.82667 | 0.93333
| Median |  |  | 0.75 | 0.89309 | 0.8 | 0.7843 | 0.8 |  | 0.9 | 0.85 | 0.9
| Std Dev. |  |  | 0.054772 | 0.079817 | 0.1169 | 0.036223 | 0.040825 |  | 0.040825 | 0.08165 | 0.10328
| Rel.Std.Dev. |  |  | 7.303% | 8.791% | 14.31% | 4.658% | 4.999% |  | 4.454% | 9.877% | 11.07%
| PDM3 |  |  | -11.05% | 7.677% | -3.146% | -7.775% | -3.146% |  | 8.713% | -1.96% | 10.69%
| Z-Score (Absolute) |  | 1.52 | 0.61 | 0.74 | 0.04 | 0.38 | 0.04 | 2.75 | 0.81 | 0.04 | 0.95
| Analytical results for Zn in OREAS 196 (Certified Value 411 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 20 |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS |  |  | 0
0
| 1 | 1 | 400 | 460.4558 | 415.5517 | 450 | 400 | 430 | 430 | 349.4767 | 444 | 420 | 361.9 | 426.2 | 379 |  |  | 1
| 1 | 2 | 400 | 454.9383 | 403.955 | 450 | 420 | 420 | 430 | 346.7266 | 449 | 419 | 362.6 | 416.2 | 366 |  |  | #N/A
| 1 | 3 | 400 | 462.6536 | 399.7288 | 450 | 420 | 410 | 430 | 336.1369 | 440 | 409 | 361.9 | 427.8 | 371 |  |  | 16
| 1 | 4 | 400 | 471.6399 | 420.9521 | 440 | 400 | 420 | 430 | 334.4902 | 439 | 420 | 360.9 | 393.4 | 375 |  |  | 411.2287
| 1 | 5 | 400 | 461.3326 | 423.1009 | 460 | 410 | 420 | 450 | 343.7744 | 442 | 420 | 365.2 | 396.2 | 380 |  |  | 72
| 1 | 6 | 400 | 460.9438 | 406.2397 | 450 | 410 | 410 | 450 | 336.2412 | 448 | 419 | 365.2 | 406 | 414
| Mean |  | 400 | 461.994 | 411.588 | 450 | 410 | 418.3333 | 436.6667 | 341.141 | 443.6667 | 417.8333 | 362.95 | 410.9667 | 380.8333
| Median |  | 400 | 461.1382 | 410.8957 | 450 | 410 | 420 | 430 | 340.0078 | 443 | 419.5 | 362.25 | 411.1 | 377
| Std Dev. |  | 0 | 5.4253 | 9.6276 | 6.3246 | 8.9443 | 7.5277 | 10.328 | 6.3388 | 4.1312 | 4.3551 | 1.8251 | 14.7927 | 17.0577
| Rel.Std.Dev. |  | 0% | 1.174% | 2.339% | 1.405% | 2.182% | 1.799% | 2.365% | 1.858% | 0.9311% | 1.042% | 0.5029% | 3.599% | 4.479%
| PDM3 |  | -2.731% | 12.34% | 0.08738% | 9.428% | -0.2988% | 1.728% | 6.186% | -17.04% | 7.888% | 1.606% | -11.74% | -0.06372% | -7.391%
| Z-Score (Absolute) |  | 0.31 | 1.36 | 0 | 1.03 | 0.04 | 0.18 | 0.67 | 1.89 | 0.86 | 0.17 | 1.31 | 0.02 | 0.83
| Analytical results for Zr in OREAS 196 (Indicative Value 9.52 ppm) | [16 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 12 | 14 | 15 | [10 empty] | ppm
|  |  | PF*MS | BF*OES/MS | PF*MS | BF*OES | PF*OES | [10 empty] | 2
2
| 1 | 1 | 9 | 9.6652 | 10.1 | < 20 | < 1 | [10 empty] | 1
| 1 | 2 | 9 | 10.0325 | 10.2 | < 20 | < 1 | [10 empty] | 8
| 1 | 3 | 9 | 9.1101 | 9.5 | < 20 | < 1 | [10 empty] | 16
| 1 | 4 | 9 | 9.2493 | 10.4 | < 20 | < 1 | [10 empty] | 9.5192
| 1 | 5 | 9 | 8.958 | 9.8 | < 20 | < 1 | [10 empty] | 24
| 1 | 6 | 9 | 10.1301 | 10.2 | < 20 | < 1
| Mean |  | 9 | 9.5242 | 10.0333
| Median |  | 9 | 9.4572 | 10.15
| Std Dev. |  | 0 | 0.49255 | 0.3266
| Rel.Std.Dev. |  | 0% | 5.172% | 3.255%
| PDM3 |  | -5.454% | 0.05289% | 5.401%
| Z-Score (Absolute) |  | 0.69 | 0 | 0.67 | 0.63 | 11.95
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C551: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C569: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C605: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C660: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C714: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C752: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C770: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C788: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C825: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C897: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C952: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 196 (Certified Value 0.563 wt.%) | [21 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 11 | 12 | 13 | 14 | 16 | 18 | 19 | 20 | 21 | 22 | 23 | 24 |  | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC |  | 3
3
| 1 | 1 | 0.55 | 0.59 | 0.56188 | 0.55094 | 0.6 | 0.63 | 0.57 | 0.55 | 0.56 | 0.54 | 0.54 | 0.53 | 0.51 | 0.6 | 0.58 | 0.57 | 0.67 | 0.59 | 0.55 |  | 1
| 1 | 2 | 0.56 | 0.6 | 0.54249 | 0.55689 | 0.61 | 0.58 | 0.56 | 0.55 | 0.557 | 0.55 | 0.54 | 0.54 | 0.54 | 0.62 | 0.56 | 0.57 | 0.65 | 0.6 | 0.562 |  | #N/A
| 1 | 3 | 0.54 | 0.59 | 0.56342 | 0.54117 | 0.61 | 0.64 | 0.56 | 0.57 | 0.563 | 0.54 | 0.56 | 0.53 | 0.52 | 0.59 | 0.57 | 0.57 | 0.66 | 0.6 | 0.558 |  | 16
| 1 | 4 | 0.57 | 0.61 | 0.57018 | 0.55489 | 0.6 | 0.63 | 0.56 | 0.55 | 0.561 | 0.53 | 0.56 | 0.54 | 0.51 | 0.61 | 0.57 | 0.57 | 0.68 | 0.6 | 0.563 |  | 0.56269
| 1 | 5 | 0.56 | 0.6 | 0.56881 | 0.55849 | 0.6 | 0.61 | 0.57 | 0.56 | 0.564 | 0.55 | 0.56 | 0.56 | 0.55 | 0.6 | 0.57 | 0.58 | 0.66 | 0.59 | 0.557 |  | 74
| 1 | 6 | 2.1 | 0.61 | 0.5611 | 0.53835 | 0.59 | 0.61 | 0.56 | 0.56 | 0.559 | 0.54 | 0.56 | 0.53 | 0.53 | 0.62 | 0.57 | 0.56 | 0.66 | 0.6 | 0.55
| Mean |  | 0.81333 | 0.6 | 0.56131 | 0.55012 | 0.60167 | 0.61667 | 0.56333 | 0.55667 | 0.56067 | 0.54167 | 0.55333 | 0.53833 | 0.52667 | 0.60667 | 0.57 | 0.57 | 0.66333 | 0.59667 | 0.55667
| Median |  | 0.56 | 0.6 | 0.56265 | 0.55291 | 0.6 | 0.62 | 0.56 | 0.555 | 0.5605 | 0.54 | 0.56 | 0.535 | 0.525 | 0.605 | 0.57 | 0.57 | 0.66 | 0.6 | 0.5575
| Std Dev. |  | 0.63042 | 0.0089443 | 0.0099366 | 0.0084597 | 0.0075277 | 0.021602 | 0.005164 | 0.008165 | 0.002582 | 0.0075277 | 0.010328 | 0.01169 | 0.01633 | 0.012111 | 0.0063246 | 0.0063246 | 0.010328 | 0.005164 | 0.0056451
| Rel.Std.Dev. |  | 77.51% | 1.491% | 1.77% | 1.538% | 1.251% | 3.503% | 0.9167% | 1.467% | 0.4605% | 1.39% | 1.866% | 2.172% | 3.101% | 1.996% | 1.11% | 1.11% | 1.557% | 0.8655% | 1.014%
| PDM3 |  | 44.54% | 6.63% | -0.2453% | -2.234% | 6.926% | 9.592% | 0.1137% | -1.071% | -0.3603% | -3.737% | -1.664% | -4.329% | -6.403% | 7.815% | 1.298% | 1.298% | 17.89% | 6.038% | -1.071%
| Z-Score (Absolute) |  | 7.78 | 1.14 | 0.06 | 0.41 | 1.19 | 1.66 | 0 | 0.21 | 0.08 | 0.67 | 0.31 | 0.78 | 1.14 | 1.35 | 0.21 | 0.21 | 3.11 | 1.04 | 0.21
| Analytical results for S in OREAS 196 (Certified Value 0.182 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 16 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.14 | 0.19 | 0.18383 | 0.19891 | 0.19 | 0.18 | 0.19 | 0.18966 | 0.17 | 0.177 | 0.17 | 0.19 | 0.14 | 0.19 | 0.22 | 0.19 | 0.19 | 0.21 | 0.17 | 0.166 | 1
| 1 | 2 | 0.16 | 0.19 | 0.18528 | 0.1988 | 0.19 | 0.18 | 0.18 | 0.18103 | 0.17 | 0.176 | 0.16 | 0.18 | 0.14 | 0.19 | 0.22 | 0.19 | 0.18 | 0.18 | 0.18 | 0.163 | #N/A
| 1 | 3 | 0.14 | 0.19 | 0.17015 | 0.19185 | 0.18 | 0.19 | 0.18 | 0.17241 | 0.17 | 0.179 | 0.16 | 0.2 | 0.15 | 0.19 | 0.22 | 0.19 | 0.18 | 0.18 | 0.18 | 0.157 | 16
| 1 | 4 | 0.16 | 0.2 | 0.18553 | 0.19347 | 0.19 | 0.18 | 0.19 | 0.18103 | 0.17 | 0.18 | 0.15 | 0.19 | 0.14 | 0.19 | 0.22 | 0.19 | 0.19 | 0.19 | 0.17 | 0.157 | 0.18228
| 1 | 5 | 0.15 | 0.2 | 0.18327 | 0.19223 | 0.2 | 0.18 | 0.19 | 0.18966 | 0.17 | 0.182 | 0.17 | 0.2 | 0.16 | 0.19 | 0.22 | 0.19 | 0.19 | 0.18 | 0.17 | 0.161 | 75
| 1 | 6 | 0.15 | 0.2 | 0.18031 | 0.19093 | 0.18 | 0.19 | 0.19 | 0.18966 | 0.17 | 0.178 | 0.16 | 0.19 | 0.14 | 0.19 | 0.22 | 0.2 | 0.19 | 0.18 | 0.17 | 0.157
| Mean |  | 0.15 | 0.195 | 0.1814 | 0.19437 | 0.18833 | 0.18333 | 0.18667 | 0.18391 | 0.17 | 0.17867 | 0.16167 | 0.19167 | 0.145 | 0.19 | 0.22 | 0.19167 | 0.18667 | 0.18667 | 0.17333 | 0.16017
| Median |  | 0.15 | 0.195 | 0.18355 | 0.19285 | 0.19 | 0.18 | 0.19 | 0.18534 | 0.17 | 0.1785 | 0.16 | 0.19 | 0.14 | 0.19 | 0.22 | 0.19 | 0.19 | 0.18 | 0.17 | 0.159
| Std Dev. |  | 0.0089443 | 0.0054772 | 0.0058184 | 0.0035714 | 0.0075277 | 0.005164 | 0.005164 | 0.0070388 | 0 | 0.0021602 | 0.0075277 | 0.0075277 | 0.0083666 | 3.0405e-17 | 0 | 0.0040825 | 0.005164 | 0.012111 | 0.005164 | 0.0038166
| Rel.Std.Dev. |  | 5.963% | 2.809% | 3.208% | 1.837% | 3.997% | 2.817% | 2.766% | 3.827% | 0% | 1.209% | 4.656% | 3.928% | 5.77% | 1.6e-14% | 0% | 2.13% | 2.766% | 6.488% | 2.979% | 2.383%
| PDM3 |  | -17.71% | 6.976% | -0.4879% | 6.627% | 3.318% | 0.5754% | 2.404% | 0.8907% | -6.739% | -1.985% | -11.31% | 5.147% | -20.45% | 4.233% | 20.69% | 5.147% | 2.404% | 2.404% | -4.91% | -12.13%
| Z-Score (Absolute) |  | 3.66 | 1.01 | 0.4 | 0.94 | 0.32 | 0.2 | 0.14 | 0.14 | 1.59 | 0.69 | 2.45 | 0.66 | 4.18 | 0.49 | 3.6 | 0.66 | 0.14 | 0.14 | 1.24 | 2.61
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Co in OREAS 196 (Indicative Value 589 ppm) | [6 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 06 | 19 | 21 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | NA | 593.2956 | - | N.A. |  | 1
| 1 | 2 | NA | 592.3515 | - | N.A. |  | 21
| 1 | 3 | NA | 590.1814 | - | N.A. |  | 16
| 1 | 4 | NA | 583.0021 | - | N.A. |  | 589.1552
| 1 | 5 | NA | 587.4344 | - | N.A. |  | 27
| 1 | 6 | NA | 588.6663 | - | N.A.
| Mean |  |  | 589.1552
| Median |  |  | 589.4239
| Std Dev. |  |  | 3.7276
| Rel.Std.Dev. |  |  | 0.6327%
| PDM3 |  |  | -2.22e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Fe in OREAS 196 (Indicative Value 17.82 wt.%) | [6 empty] | Indicative
Fe | Round | Replicate | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 06 | 19 | 21 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS |  | 2
2
| 1 | 1 | NA | 17.8205 | - | N.A. |  | 1
| 1 | 2 | NA | 17.9887 | - | N.A. |  | 22
| 1 | 3 | NA | 17.7932 | - | N.A. |  | 16
| 1 | 4 | NA | 17.6877 | - | N.A. |  | 17.8223
| 1 | 5 | NA | 17.8031 | - | N.A. |  | 28
| 1 | 6 | NA | 17.8406 | - | N.A.
| Mean |  |  | 17.8223
| Median |  |  | 17.8118
| Std Dev. |  |  | 0.097304
| Rel.Std.Dev. |  |  | 0.546%
| PDM3 |  |  | 1.776e-13%
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Mg in OREAS 196 (Indicative Value 7.24 wt.%) | [6 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 06 | 19 | 21 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS |  | 2
2
| 1 | 1 | NA | 7.26 | - | N.A. |  | 1
| 1 | 2 | NA | 7.2959 | - | N.A. |  | 21
| 1 | 3 | NA | 7.2607 | - | N.A. |  | 16
| 1 | 4 | NA | 7.1799 | - | N.A. |  | 7.2439
| 1 | 5 | NA | 7.2531 | - | N.A. |  | 27
| 1 | 6 | NA | 7.2135 | - | N.A.
| Mean |  |  | 7.2439
| Median |  |  | 7.2566
| Std Dev. |  |  | 0.040859
| Rel.Std.Dev. |  |  | 0.5641%
| PDM3 |  |  | 4.441e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Ni in OREAS 196 (Indicative Value 2.05 wt.%) | [6 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 06 | 15 | 19 | 21 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 2
2
| 1 | 1 | NA | 2.0691 | 2.04 | - | N.A. | 1
| 1 | 2 | NA | 2.0602 | 2.04 | - | N.A. | 22
| 1 | 3 | NA | 2.0599 | 2.04 | - | N.A. | 16
| 1 | 4 | NA | 2.0321 | 2.04 | - | N.A. | 2.0473
| 1 | 5 | NA | 2.0538 | 2.04 | - | N.A. | 28
| 1 | 6 | NA | 2.0528 | 2.04 | - | N.A.
| Mean |  |  | 2.0547 | 2.04
| Median |  |  | 2.0569 | 2.04
| Std Dev. |  |  | 0.012485 | 4.8648e-16
| Rel.Std.Dev. |  |  | 0.6076% | 2.385e-14%
| PDM3 |  |  | 0.3581% | -0.3581%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67 | NA | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.