## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 230b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.38099 | 0.37552 | 0.38646 | 0.37718 | 0.3848
PhotonAssay
Au, Gold (ppm) | 0.37365 | 0.36698 | 0.38033 | 0.36991 | 0.37739
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.36806 | 0.357 | 0.37913 | 0.36403 | 0.3721
Cyanide Leach
Au, Gold (ppm) | 0.36314 | 0.35156 | 0.37471 | 0.36173 | 0.36454
4-Acid Digestion
Ag, Silver (ppm) | 0.19483 | 0.17119 | 0.21848 | 0.17029 | 0.21938
Al, Aluminium (wt.%) | 6.8164 | 6.6411 | 6.9916 | 6.6954 | 6.9373
As, Arsenic (ppm) | 26.8227 | 25.686 | 27.9595 | 25.9092 | 27.7362
Ba, Barium (ppm) | 94.1885 | 90.2656 | 98.1115 | 92.1396 | 96.2375
Be, Beryllium (ppm) | 0.37266 | 0.35058 | 0.39474 | 0.35831 | 0.387
Bi, Bismuth (ppm) | 0.23144 | 0.21373 | 0.24914 | 0.21532 | 0.24755
Ca, Calcium (wt.%) | 7.2202 | 7.0361 | 7.4042 | 7.1451 | 7.2952
Cd, Cadmium (ppm) | 0.30789 | 0.2728 | 0.34299 | 0.27941 | 0.33638
Ce, Cerium (ppm) | 11.06 | 10.4713 | 11.6487 | 10.6558 | 11.4641
Co, Cobalt (ppm) | 44.397 | 43.0037 | 45.7904 | 43.4688 | 45.3253
Cr, Chromium (ppm) | 86.2362 | 78.6915 | 93.7809 | 83.2572 | 89.2152
Cs, Caesium (ppm) | 0.41853 | 0.39437 | 0.44268 | 0.39684 | 0.44021
Cu, Copper (ppm) | 179.7329 | 174.4963 | 184.9695 | 176.6689 | 182.7969
Dy, Dysprosium (ppm) | 4.2489 | 3.7511 | 4.7467 | 4.1286 | 4.3692
Er, Erbium (ppm) | 2.6669 | 2.2893 | 3.0446 | 2.5186 | 2.8153
Eu, Europium (ppm) | 0.99936 | 0.89332 | 1.1054 | 0.92515 | 1.0736
Fe, Iron (wt.%) | 8.9645 | 8.7478 | 9.1812 | 8.8694 | 9.0596
Ga, Gallium (ppm) | 15.8493 | 15.1296 | 16.569 | 15.3406 | 16.358
Gd, Gadolinium (ppm) | 3.6156 | 3.222 | 4.0092 | 3.4482 | 3.783
Hf, Hafnium (ppm) | 1.784 | 1.6116 | 1.9564 | 1.675 | 1.8929
Ho, Holmium (ppm) | 0.91221 | 0.79184 | 1.0326 | 0.86453 | 0.9599
In, Indium (ppm) | 0.080185 | 0.0729 | 0.087471 | 0.07439 | 0.08598
K, Potassium (wt.%) | 0.18033 | 0.1726 | 0.18806 | 0.17478 | 0.18588
La, Lanthanum (ppm) | 4.3494 | 4.1207 | 4.5782 | 4.1899 | 4.509
Li, Lithium (ppm) | 10.1106 | 9.6399 | 10.5813 | 9.883 | 10.3383
Lu, Lutetium (ppm) | 0.37419 | 0.34072 | 0.40767 | 0.34501 | 0.40337
Mg, Magnesium (wt.%) | 3.8181 | 3.7232 | 3.9129 | 3.7577 | 3.8784
Mn, Manganese (wt.%) | 0.1506 | 0.14544 | 0.15577 | 0.14792 | 0.15329
Mo, Molybdenum (ppm) | 0.56889 | 0.42496 | 0.71281 | 0.48789 | 0.64988
Na, Sodium (wt.%) | 1.9961 | 1.9505 | 2.0418 | 1.9679 | 2.0243
Nb, Niobium (ppm) | 3.7505 | 3.5505 | 3.9506 | 3.603 | 3.898
Nd, Neodymium (ppm) | 8.6714 | 7.8361 | 9.5068 | 8.4286 | 8.9143
Ni, Nickel (ppm) | 68.2634 | 65.3977 | 71.129 | 66.6494 | 69.8774
P, Phosphorus (wt.%) | 0.051655 | 0.04993 | 0.053379 | 0.050528 | 0.052781
Pb, Lead (ppm) | 9.5883 | 8.6208 | 10.5558 | 9.0828 | 10.0938
Pr, Praseodymium (ppm) | 1.6503 | 1.5385 | 1.7621 | 1.597 | 1.7036
Rb, Rubidium (ppm) | 4.6408 | 4.3555 | 4.9261 | 4.4306 | 4.851
Re, Rhenium (ppm) | 0.0024542 | 0.0012051 | 0.0037033 | IND | IND
S, Sulphur (wt.%) | 0.2716 | 0.26295 | 0.28025 | 0.26548 | 0.27771
Sc, Scandium (ppm) | 43.1179 | 41.7112 | 44.5247 | 42.1298 | 44.1061
Se, Selenium (ppm) | 0.82406 | 0.61988 | 1.0282 | IND | IND
Sm, Samarium (ppm) | 2.7332 | 2.552 | 2.9143 | 2.5803 | 2.886
Sn, Tin (ppm) | 0.80441 | 0.70206 | 0.90675 | 0.67372 | 0.93509
Sr, Strontium (ppm) | 119.7834 | 115.7407 | 123.8262 | 117.0536 | 122.5133
Ta, Tantalum (ppm) | 0.23729 | 0.21842 | 0.25616 | 0.21934 | 0.25525
Tb, Terbium (ppm) | 0.61178 | 0.57242 | 0.65113 | 0.5873 | 0.63625
Te, Tellurium (ppm) | 0.11556 | 0.074751 | 0.15636 | IND | IND
Th, Thorium (ppm) | 0.52039 | 0.47907 | 0.5617 | 0.49716 | 0.54362
Ti, Titanium (wt.%) | 0.69656 | 0.67711 | 0.71601 | 0.68333 | 0.70978
Tl, Thallium (ppm) | 0.078315 | 0.070219 | 0.086411 | IND | IND
Tm, Thulium (ppm) | 0.37867 | 0.33981 | 0.41753 | 0.32999 | 0.42735
U, Uranium (ppm) | 0.17979 | 0.16612 | 0.19345 | IND | IND
V, Vanadium (ppm) | 314.3412 | 303.9462 | 324.7362 | 307.9165 | 320.7659
W, Tungsten (ppm) | 9.1405 | 8.8094 | 9.4716 | 8.8302 | 9.4508
Y, Yttrium (ppm) | 23.3568 | 22.4193 | 24.2944 | 22.8186 | 23.895
Yb, Ytterbium (ppm) | 2.5266 | 2.2879 | 2.7653 | 2.4482 | 2.605
Zn, Zinc (ppm) | 109.097 | 104.374 | 113.82 | 106.8626 | 111.3314
Zr, Zirconium (ppm) | 55.8758 | 50.7189 | 61.0326 | 53.7069 | 58.0447
Aqua Regia Digestion
Ag, Silver (ppm) | 0.18068 | 0.16412 | 0.19723 | 0.16276 | 0.19859
Al, Aluminium (wt.%) | 3.3989 | 3.24 | 3.5578 | 3.3229 | 3.4748
As, Arsenic (ppm) | 25.6842 | 24.0572 | 27.3111 | 24.769 | 26.5993
Ba, Barium (ppm) | 19.7329 | 18.9701 | 20.4957 | 18.944 | 20.5218
Be, Beryllium (ppm) | 0.15086 | 0.11957 | 0.18216 | IND | IND
Bi, Bismuth (ppm) | 0.23913 | 0.22344 | 0.25481 | 0.22489 | 0.25336
Ca, Calcium (wt.%) | 3.112 | 2.9033 | 3.3207 | 3.0359 | 3.1881
Cd, Cadmium (ppm) | 0.26691 | 0.2443 | 0.28952 | 0.24544 | 0.28838
Ce, Cerium (ppm) | 6.1488 | 5.6738 | 6.6238 | 5.9451 | 6.3525
Co, Cobalt (ppm) | 32.5505 | 30.9095 | 34.1915 | 31.717 | 33.3841
Cr, Chromium (ppm) | 21.5785 | 20.2856 | 22.8714 | 20.6252 | 22.5319
Cs, Caesium (ppm) | 0.33828 | 0.31805 | 0.35852 | 0.32 | 0.35656
Cu, Copper (ppm) | 178.702 | 171.0911 | 186.313 | 174.4139 | 182.9902
Fe, Iron (wt.%) | 6.3768 | 6.1421 | 6.6115 | 6.2698 | 6.4838
Ga, Gallium (ppm) | 11.0779 | 10.2369 | 11.9189 | 10.6963 | 11.4596
Ge, Germanium (ppm) | 0.17349 | 0.12603 | 0.22094 | IND | IND
Hf, Hafnium (ppm) | 0.50893 | 0.46617 | 0.55168 | 0.4907 | 0.52716
In, Indium (ppm) | 0.030191 | 0.027674 | 0.032708 | 0.028471 | 0.03191
K, Potassium (wt.%) | 0.050052 | 0.049253 | 0.050852 | 0.047623 | 0.052481
La, Lanthanum (ppm) | 2.3379 | 2.1674 | 2.5084 | 2.232 | 2.4438
Li, Lithium (ppm) | 9.2157 | 8.8763 | 9.555 | 8.9538 | 9.4775
Mg, Magnesium (wt.%) | 1.9593 | 1.9036 | 2.0149 | 1.9155 | 2.003
Mn, Manganese (wt.%) | 0.088459 | 0.084769 | 0.09215 | 0.087186 | 0.089732
Mo, Molybdenum (ppm) | 0.51429 | 0.46893 | 0.55966 | 0.47838 | 0.5502
Na, Sodium (wt.%) | 0.10814 | 0.096176 | 0.1201 | 0.097795 | 0.11849
Ni, Nickel (ppm) | 44.4146 | 41.8409 | 46.9883 | 43.189 | 45.6402
P, Phosphorus (wt.%) | 0.047581 | 0.045456 | 0.049707 | 0.046367 | 0.048795
Pb, Lead (ppm) | 8.4112 | 7.9633 | 8.8592 | 8.1714 | 8.6511
Rb, Rubidium (ppm) | 2.2428 | 2.083 | 2.4026 | 2.1391 | 2.3465
S, Sulphur (wt.%) | 0.26634 | 0.25305 | 0.27964 | 0.25517 | 0.27752
Sb, Antimony (ppm) | 0.51635 | 0.36702 | 0.66568 | 0.44816 | 0.58454
Sc, Scandium (ppm) | 6.0064 | 5.0997 | 6.913 | 5.7713 | 6.2414
Se, Selenium (ppm) | 0.58389 | 0.45381 | 0.71397 | IND | IND
Sn, Tin (ppm) | 0.46344 | 0.37372 | 0.55316 | 0.38361 | 0.54328
Sr, Strontium (ppm) | 23.487 | 21.0613 | 25.9128 | 22.8085 | 24.1656
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Te, Tellurium (ppm) | 0.10263 | 0.085209 | 0.12006 | IND | IND
Th, Thorium (ppm) | 0.35382 | 0.32668 | 0.38096 | IND | IND
Ti, Titanium (wt.%) | 0.35895 | 0.31302 | 0.40488 | 0.34672 | 0.37119
Tl, Thallium (ppm) | 0.045241 | 0.035549 | 0.054933 | IND | IND
U, Uranium (ppm) | 0.10546 | 0.092209 | 0.11871 | IND | IND
V, Vanadium (ppm) | 168.6942 | 156.8477 | 180.5407 | 165.4968 | 171.8916
W, Tungsten (ppm) | 6.591 | 5.9424 | 7.2396 | 6.3798 | 6.8021
Y, Yttrium (ppm) | 12.1372 | 10.9827 | 13.2917 | 11.8603 | 12.4141
Zn, Zinc (ppm) | 95.7961 | 92.3553 | 99.2369 | 93.2173 | 98.3749
Zr, Zirconium (ppm) | 18.1371 | 15.7289 | 20.5453 | 17.5016 | 18.7726
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 230b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Ge | ppm | 0.12167 | Hg | ppm | 0.068333 | Sb | ppm | 1.4365
Aqua Regia Digestion
B | ppm | 13.5038 | Ho | ppm | 0.51494 | Sm | ppm | 1.5122
Dy | ppm | 2.4568 | Lu | ppm | 0.17702 | Tb | ppm | 0.40605
Er | ppm | 1.5558 | Nb | ppm | 0.16266 | Tm | ppm | 0.23753
Eu | ppm | 0.52586 | Nd | ppm | 4.7443 | Yb | ppm | 1.2003
Gd | ppm | 1.9206 | Pr | ppm | 0.98523
Hg | ppm | 0.030545 | Re | ppm | 0.002125
Borate Fusion XRF
Al2O3 | wt.% | 13.11 | MgO | wt.% | 6.555 | SiO2 | wt.% | 48.25
CaO | wt.% | 10.38 | MnO | wt.% | 0.2 | TiO2 | wt.% | 1.2
Fe2O3 | wt.% | 13.255 | P2O5 | wt.% | 0.1135
K2O | wt.% | 0.2055 | S | wt.% | 0.27811
Thermogravimetry
LOI1000 | wt.% | 3.7
Infrared Combustion
C | wt.% | 0.375 | S | wt.% | 0.305
Laser Ablation ICP-MS
Ag | ppm | 0.15 | Hf | ppm | 1.935 | Sn | ppm | 0.8
As | ppm | 24.7 | Ho | ppm | 0.93 | Sr | ppm | 112
Ba | ppm | 91.75 | In | ppm | 0.05 | Ta | ppm | 0.65
Be | ppm | 0.6 | La | ppm | 4.365 | Tb | ppm | 0.63
Bi | ppm | 0.22 | Lu | ppm | 0.375 | Te | ppm | < 0.2
Cd | ppm | 0.2 | Mn | wt.% | 0.1615 | Th | ppm | 0.465
Ce | ppm | 10.7 | Mo | ppm | 0.6 | Ti | wt.% | 0.708
Co | ppm | 47.1 | Nb | ppm | 3.7 | Tl | ppm | < 0.2
Cr | ppm | 100 | Nd | ppm | 8.25 | Tm | ppm | 0.375
Cs | ppm | 0.455 | Ni | ppm | 74 | U | ppm | 0.13
Cu | ppm | 186 | Pb | ppm | 8 | V | ppm | 337
Dy | ppm | 4.03 | Pr | ppm | 1.705 | W | ppm | 9.225
Er | ppm | 2.67 | Rb | ppm | 4.6 | Y | ppm | 23.25
Eu | ppm | 1 | Re | ppm | < 0.01 | Yb | ppm | 2.68
Ga | ppm | 15.25 | Sb | ppm | 1.6 | Zn | ppm | 107.5
Gd | ppm | 3.545 | Sc | ppm | 44.15 | Zr | ppm | 66.5
Ge | ppm | 1.125 | Sm | ppm | 2.78
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 230b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.38099 | 0.012131 | 0.35673 | 0.40525 | 0.3446 | 0.41739 | 3.184% | 6.368% | 9.552% | 0.36194 | 0.40004
PhotonAssay
Au, ppm | 0.37365 | 0.022023 | 0.32961 | 0.4177 | 0.30758 | 0.43972 | 5.894% | 11.79% | 17.68% | 0.35497 | 0.39234
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.36806 | 0.022664 | 0.32274 | 0.41339 | 0.30007 | 0.43605 | 6.158% | 12.32% | 18.47% | 0.34966 | 0.38647
Cyanide Leach
Au, ppm | 0.36314 | 0.026156 | 0.31083 | 0.41545 | 0.28467 | 0.44161 | 7.203% | 14.41% | 21.61% | 0.34498 | 0.38129
4-Acid Digestion
Ag, ppm | 0.19483 | 0.016627 | 0.16158 | 0.22809 | 0.14495 | 0.24472 | 8.534% | 17.07% | 25.6% | 0.18509 | 0.20458
Al, wt.% | 6.8164 | 0.18505 | 6.4463 | 7.1865 | 6.2612 | 7.3715 | 2.715% | 5.43% | 8.144% | 6.4755 | 7.1572
As, ppm | 26.8227 | 1.2846 | 24.2535 | 29.392 | 22.9689 | 30.6766 | 4.789% | 9.579% | 14.37% | 25.4816 | 28.1639
Ba, ppm | 94.1885 | 3.9477 | 86.2932 | 102.0839 | 82.3455 | 106.0316 | 4.191% | 8.383% | 12.57% | 89.4791 | 98.898
Be, ppm | 0.37266 | 0.023693 | 0.32527 | 0.42004 | 0.30158 | 0.44374 | 6.358% | 12.72% | 19.07% | 0.35402 | 0.39129
Bi, ppm | 0.23144 | 0.019285 | 0.19287 | 0.27001 | 0.17358 | 0.28929 | 8.333% | 16.67% | 25% | 0.21987 | 0.24301
Ca, wt.% | 7.2202 | 0.14672 | 6.9267 | 7.5136 | 6.78 | 7.6603 | 2.032% | 4.064% | 6.096% | 6.8591 | 7.5812
Cd, ppm | 0.30789 | 0.025191 | 0.25751 | 0.35828 | 0.23232 | 0.38347 | 8.182% | 16.36% | 24.54% | 0.2925 | 0.32329
Ce, ppm | 11.06 | 0.66479 | 9.7304 | 12.3896 | 9.0656 | 13.0543 | 6.011% | 12.02% | 18.03% | 10.507 | 11.613
Co, ppm | 44.397 | 1.9857 | 40.4257 | 48.3684 | 38.44 | 50.3541 | 4.473% | 8.945% | 13.42% | 42.1772 | 46.6169
Cr, ppm | 86.2362 | 11.6684 | 62.8993 | 109.5731 | 51.2309 | 121.2415 | 13.53% | 27.06% | 40.59% | 81.9244 | 90.548
Cs, ppm | 0.41853 | 0.0284 | 0.36173 | 0.47533 | 0.33333 | 0.50373 | 6.786% | 13.57% | 20.36% | 0.3976 | 0.43945
Cu, ppm | 179.7329 | 6.6029 | 166.5271 | 192.9387 | 159.9241 | 199.5417 | 3.674% | 7.347% | 11.02% | 170.7463 | 188.7195
Dy, ppm | 4.2489 | 0.41539 | 3.4181 | 5.0797 | 3.0027 | 5.4951 | 9.776% | 19.55% | 29.33% | 4.0365 | 4.4613
Er, ppm | 2.6669 | 0.16114 | 2.3447 | 2.9892 | 2.1835 | 3.1504 | 6.042% | 12.08% | 18.13% | 2.5336 | 2.8003
Eu, ppm | 0.99936 | 0.071753 | 0.85586 | 1.1429 | 0.7841 | 1.2146 | 7.18% | 14.36% | 21.54% | 0.94939 | 1.0493
Fe, wt.% | 8.9645 | 0.2193 | 8.5259 | 9.4031 | 8.3066 | 9.6224 | 2.446% | 4.893% | 7.339% | 8.5163 | 9.4127
Ga, ppm | 15.8493 | 0.74529 | 14.3587 | 17.3399 | 13.6134 | 18.0852 | 4.702% | 9.405% | 14.11% | 15.0568 | 16.6418
Gd, ppm | 3.6156 | 0.17153 | 3.2725 | 3.9587 | 3.101 | 4.1302 | 4.744% | 9.488% | 14.23% | 3.4348 | 3.7964
Hf, ppm | 1.784 | 0.16184 | 1.4603 | 2.1077 | 1.2984 | 2.2695 | 9.072% | 18.14% | 27.22% | 1.6948 | 1.8732
Ho, ppm | 0.91221 | 0.080311 | 0.75159 | 1.0728 | 0.67128 | 1.1531 | 8.804% | 17.61% | 26.41% | 0.8666 | 0.95783
In, ppm | 0.080185 | 0.0036913 | 0.072803 | 0.087568 | 0.069111 | 0.091259 | 4.603% | 9.207% | 13.81% | 0.076176 | 0.084194
K, wt.% | 0.18033 | 0.009454 | 0.16143 | 0.19924 | 0.15197 | 0.2087 | 5.242% | 10.48% | 15.73% | 0.17132 | 0.18935
La, ppm | 4.3494 | 0.24682 | 3.8558 | 4.8431 | 3.609 | 5.0899 | 5.675% | 11.35% | 17.02% | 4.132 | 4.5669
Li, ppm | 10.1106 | 0.56858 | 8.9735 | 11.2478 | 8.4049 | 11.8164 | 5.624% | 11.25% | 16.87% | 9.6051 | 10.6162
Lu, ppm | 0.37419 | 0.027254 | 0.31968 | 0.4287 | 0.29243 | 0.45595 | 7.283% | 14.57% | 21.85% | 0.35548 | 0.3929
Mg, wt.% | 3.8181 | 0.13179 | 3.5545 | 4.0817 | 3.4227 | 4.2134 | 3.452% | 6.904% | 10.36% | 3.6272 | 4.009
Mn, wt.% | 0.1506 | 0.0078055 | 0.13499 | 0.16622 | 0.12719 | 0.17402 | 5.183% | 10.37% | 15.55% | 0.14307 | 0.15813
Mo, ppm | 0.56889 | 0.11421 | 0.34046 | 0.79731 | 0.22625 | 0.91153 | 20.08% | 40.15% | 60.23% | 0.54044 | 0.59733
Na, wt.% | 1.9961 | 0.0512 | 1.8937 | 2.0985 | 1.8425 | 2.1497 | 2.565% | 5.13% | 7.695% | 1.8963 | 2.0959
Nb, ppm | 3.7505 | 0.17182 | 3.4069 | 4.0942 | 3.2351 | 4.266 | 4.581% | 9.163% | 13.74% | 3.563 | 3.938
Nd, ppm | 8.6714 | 0.56893 | 7.5336 | 9.8093 | 6.9647 | 10.3782 | 6.561% | 13.12% | 19.68% | 8.2379 | 9.105
Ni, ppm | 68.2634 | 4.051 | 60.1614 | 76.3653 | 56.1105 | 80.4163 | 5.934% | 11.87% | 17.8% | 64.8502 | 71.6765
P, wt.% | 0.051655 | 0.0012265 | 0.049202 | 0.054108 | 0.047975 | 0.055334 | 2.374% | 4.749% | 7.123% | 0.049072 | 0.054237
Pb, ppm | 9.5883 | 1.2873 | 7.0137 | 12.1629 | 5.7264 | 13.4502 | 13.43% | 26.85% | 40.28% | 9.1089 | 10.0677
Pr, ppm | 1.6503 | 0.054738 | 1.5408 | 1.7598 | 1.4861 | 1.8145 | 3.317% | 6.634% | 9.951% | 1.5678 | 1.7328
Rb, ppm | 4.6408 | 0.31667 | 4.0075 | 5.2741 | 3.6908 | 5.5908 | 6.824% | 13.65% | 20.47% | 4.4088 | 4.8728
Re, ppm | 0.0024542 | 0.00059058 | 0.001273 | 0.0036353 | 0.00068244 | 0.0042259 | 24.06% | 48.13% | 72.19% | 0.0023315 | 0.0025769
S, wt.% | 0.2716 | 0.0078323 | 0.25593 | 0.28726 | 0.2481 | 0.29509 | 2.884% | 5.768% | 8.651% | 0.25802 | 0.28518
Sc, ppm | 43.1179 | 1.394 | 40.33 | 45.9058 | 38.9361 | 47.2998 | 3.233% | 6.466% | 9.699% | 40.962 | 45.2738
Se, ppm | 0.82406 | 0.18547 | 0.45311 | 1.195 | 0.26764 | 1.3805 | 22.51% | 45.01% | 67.52% | 0.78285 | 0.86526
Sm, ppm | 2.7332 | 0.14629 | 2.4406 | 3.0257 | 2.2943 | 3.172 | 5.352% | 10.7% | 16.06% | 2.5965 | 2.8698
Sn, ppm | 0.80441 | 0.068189 | 0.66803 | 0.94078 | 0.59984 | 1.009 | 8.477% | 16.95% | 25.43% | 0.76419 | 0.84463
Sr, ppm | 119.7834 | 5.0306 | 109.7223 | 129.8445 | 104.6918 | 134.8751 | 4.2% | 8.399% | 12.6% | 113.7943 | 125.7726
Ta, ppm | 0.23729 | 0.012332 | 0.21263 | 0.26196 | 0.2003 | 0.27429 | 5.197% | 10.39% | 15.59% | 0.22543 | 0.24916
Tb, ppm | 0.61178 | 0.019336 | 0.5731 | 0.65045 | 0.55377 | 0.66979 | 3.161% | 6.321% | 9.482% | 0.58119 | 0.64237
Te, ppm | 0.11556 | 0.01889 | 0.077776 | 0.15334 | 0.058886 | 0.17223 | 16.35% | 32.69% | 49.04% | 0.10978 | 0.12133
Th, ppm | 0.52039 | 0.047141 | 0.4261 | 0.61467 | 0.37896 | 0.66181 | 9.059% | 18.12% | 27.18% | 0.49437 | 0.5464
Ti, wt.% | 0.69656 | 0.021075 | 0.65441 | 0.73871 | 0.63333 | 0.75978 | 3.026% | 6.051% | 9.077% | 0.66173 | 0.73139
Tl, ppm | 0.078315 | 0.0069205 | 0.064474 | 0.092156 | 0.057554 | 0.099077 | 8.837% | 17.67% | 26.51% | 0.074399 | 0.082231
Tm, ppm | 0.37867 | 0.024283 | 0.3301 | 0.42724 | 0.30582 | 0.45152 | 6.413% | 12.83% | 19.24% | 0.35974 | 0.3976
U, ppm | 0.17979 | 0.028267 | 0.12325 | 0.23632 | 0.094987 | 0.26459 | 15.72% | 31.44% | 47.17% | 0.1708 | 0.18878
V, ppm | 314.3412 | 11.9621 | 290.4169 | 338.2655 | 278.4548 | 350.2276 | 3.805% | 7.611% | 11.42% | 298.6241 | 330.0583
W, ppm | 9.1405 | 0.30286 | 8.5348 | 9.7462 | 8.2319 | 10.0491 | 3.313% | 6.627% | 9.94% | 8.6835 | 9.5975
Y, ppm | 23.3568 | 1.0431 | 21.2707 | 25.4429 | 20.2276 | 26.486 | 4.466% | 8.932% | 13.4% | 22.189 | 24.5246
Yb, ppm | 2.5266 | 0.14441 | 2.2378 | 2.8154 | 2.0934 | 2.9598 | 5.715% | 11.43% | 17.15% | 2.4003 | 2.6529
Zn, ppm | 109.097 | 5.9625 | 97.1721 | 121.0219 | 91.2096 | 126.9844 | 5.465% | 10.93% | 16.4% | 103.6421 | 114.5518
Zr, ppm | 55.8758 | 4.4474 | 46.981 | 64.7705 | 42.5337 | 69.2179 | 7.959% | 15.92% | 23.88% | 53.082 | 58.6696
Aqua Regia Digestion
Ag, ppm | 0.18068 | 0.015673 | 0.14933 | 0.21202 | 0.13366 | 0.2277 | 8.675% | 17.35% | 26.02% | 0.17164 | 0.18971
Al, wt.% | 3.3989 | 0.22147 | 2.9559 | 3.8418 | 2.7345 | 4.0633 | 6.516% | 13.03% | 19.55% | 3.2289 | 3.5688
As, ppm | 25.6842 | 1.3989 | 22.8863 | 28.482 | 21.4874 | 29.8809 | 5.447% | 10.89% | 16.34% | 24.4 | 26.9684
Ba, ppm | 19.7329 | 0.72317 | 18.2866 | 21.1792 | 17.5634 | 21.9024 | 3.665% | 7.33% | 10.99% | 18.7463 | 20.7195
Be, ppm | 0.15086 | 0.033453 | 0.083957 | 0.21777 | 0.050504 | 0.25122 | 22.17% | 44.35% | 66.52% | 0.14332 | 0.15841
Bi, ppm | 0.23913 | 0.013787 | 0.21155 | 0.2667 | 0.19776 | 0.28049 | 5.766% | 11.53% | 17.3% | 0.22717 | 0.25108
Ca, wt.% | 3.112 | 0.32814 | 2.4557 | 3.7683 | 2.1276 | 4.0964 | 10.54% | 21.09% | 31.63% | 2.9564 | 3.2676
Cd, ppm | 0.26691 | 0.012833 | 0.24124 | 0.29258 | 0.22841 | 0.30541 | 4.808% | 9.616% | 14.42% | 0.25357 | 0.28026
Ce, ppm | 6.1488 | 0.63226 | 4.8843 | 7.4133 | 4.252 | 8.0456 | 10.28% | 20.57% | 30.85% | 5.8413 | 6.4562
Co, ppm | 32.5505 | 1.4173 | 29.7159 | 35.3851 | 28.2987 | 36.8024 | 4.354% | 8.708% | 13.06% | 30.923 | 34.1781
Cr, ppm | 21.5785 | 1.2422 | 19.094 | 24.063 | 17.8518 | 25.3053 | 5.757% | 11.51% | 17.27% | 20.4996 | 22.6575
Cs, ppm | 0.33828 | 0.017555 | 0.30317 | 0.37339 | 0.28562 | 0.39095 | 5.189% | 10.38% | 15.57% | 0.32137 | 0.3552
Cu, ppm | 178.702 | 5.7099 | 167.2823 | 190.1218 | 161.5724 | 195.8317 | 3.195% | 6.39% | 9.586% | 169.7669 | 187.6371
Fe, wt.% | 6.3768 | 0.19695 | 5.9829 | 6.7707 | 5.7859 | 6.9677 | 3.089% | 6.177% | 9.266% | 6.058 | 6.6956
Ga, ppm | 11.0779 | 1.0045 | 9.0689 | 13.0869 | 8.0644 | 14.0914 | 9.068% | 18.14% | 27.2% | 10.524 | 11.6318
Ge, ppm | 0.17349 | 0.044422 | 0.084643 | 0.26233 | 0.04022 | 0.30675 | 25.61% | 51.21% | 76.82% | 0.16481 | 0.18216
Hf, ppm | 0.50893 | 0.045109 | 0.41871 | 0.59915 | 0.3736 | 0.64425 | 8.863% | 17.73% | 26.59% | 0.48348 | 0.53437
In, ppm | 0.030191 | 0.0025468 | 0.025097 | 0.035284 | 0.02255 | 0.037831 | 8.436% | 16.87% | 25.31% | 0.028681 | 0.0317
K, wt.% | 0.050052 | 0.00040447 | 0.049243 | 0.050861 | 0.048839 | 0.051266 | 0.8081% | 1.616% | 2.424% | 0.04755 | 0.052555
La, ppm | 2.3379 | 0.16824 | 2.0014 | 2.6744 | 1.8332 | 2.8426 | 7.196% | 14.39% | 21.59% | 2.221 | 2.4548
Li, ppm | 9.2157 | 0.29587 | 8.6239 | 9.8074 | 8.3281 | 10.1033 | 3.21% | 6.421% | 9.631% | 8.7549 | 9.6764
Mg, wt.% | 1.9593 | 0.047265 | 1.8647 | 2.0538 | 1.8175 | 2.1011 | 2.412% | 4.825% | 7.237% | 1.8613 | 2.0572
Mn, wt.% | 0.088459 | 0.00399 | 0.080479 | 0.096439 | 0.076489 | 0.10043 | 4.511% | 9.021% | 13.53% | 0.084036 | 0.092882
Mo, ppm | 0.51429 | 0.027883 | 0.45852 | 0.57006 | 0.43064 | 0.59794 | 5.422% | 10.84% | 16.27% | 0.48858 | 0.54001
Na, wt.% | 0.10814 | 0.013324 | 0.081492 | 0.13479 | 0.068168 | 0.14811 | 12.32% | 24.64% | 36.96% | 0.10273 | 0.11355
Ni, ppm | 44.4146 | 2.3986 | 39.6175 | 49.2117 | 37.2189 | 51.6103 | 5.4% | 10.8% | 16.2% | 42.1939 | 46.6353
P, wt.% | 0.047581 | 0.002658 | 0.042265 | 0.052897 | 0.039607 | 0.055555 | 5.586% | 11.17% | 16.76% | 0.045202 | 0.04996
Pb, ppm | 8.4112 | 0.47753 | 7.4562 | 9.3663 | 6.9786 | 9.8438 | 5.677% | 11.35% | 17.03% | 7.9907 | 8.8318
Rb, ppm | 2.2428 | 0.17279 | 1.8972 | 2.5884 | 1.7244 | 2.7612 | 7.704% | 15.41% | 23.11% | 2.1307 | 2.3549
S, wt.% | 0.26634 | 0.01268 | 0.24098 | 0.2917 | 0.2283 | 0.30438 | 4.761% | 9.522% | 14.28% | 0.25303 | 0.27966
Sb, ppm | 0.51635 | 0.12406 | 0.26822 | 0.76448 | 0.14416 | 0.88854 | 24.03% | 48.05% | 72.08% | 0.49053 | 0.54217
Sc, ppm | 6.0064 | 1.3317 | 3.3429 | 8.6698 | 2.0112 | 10.0015 | 22.17% | 44.34% | 66.52% | 5.7061 | 6.3067
Se, ppm | 0.58389 | 0.10471 | 0.37448 | 0.7933 | 0.26977 | 0.89801 | 17.93% | 35.86% | 53.8% | 0.55469 | 0.61308
Sn, ppm | 0.46344 | 0.1175 | 0.22844 | 0.69845 | 0.11093 | 0.81595 | 25.35% | 50.71% | 76.06% | 0.44027 | 0.48661
Sr, ppm | 23.487 | 3.5082 | 16.4706 | 30.5034 | 12.9624 | 34.0116 | 14.94% | 29.87% | 44.81% | 22.3127 | 24.6614
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Te, ppm | 0.10263 | 0.012623 | 0.077388 | 0.12788 | 0.064765 | 0.1405 | 12.3% | 24.6% | 36.9% | 0.097502 | 0.10777
Th, ppm | 0.35382 | 0.046997 | 0.25983 | 0.44781 | 0.21283 | 0.49481 | 13.28% | 26.57% | 39.85% | 0.33613 | 0.37151
Ti, wt.% | 0.35895 | 0.06945 | 0.22005 | 0.49785 | 0.1506 | 0.5673 | 19.35% | 38.7% | 58.04% | 0.34101 | 0.3769
Tl, ppm | 0.045241 | 0.0049865 | 0.035268 | 0.055214 | 0.030281 | 0.0602 | 11.02% | 22.04% | 33.07% | 0.042979 | 0.047503
U, ppm | 0.10546 | 0.012776 | 0.079908 | 0.13101 | 0.067132 | 0.14379 | 12.11% | 24.23% | 36.34% | 0.10019 | 0.11073
V, ppm | 168.6942 | 15.6104 | 137.4734 | 199.915 | 121.8629 | 215.5255 | 9.254% | 18.51% | 27.76% | 160.2595 | 177.1289
W, ppm | 6.591 | 0.88834 | 4.8143 | 8.3677 | 3.926 | 9.256 | 13.48% | 26.96% | 40.43% | 6.2614 | 6.9205
Y, ppm | 12.1372 | 1.6595 | 8.8182 | 15.4562 | 7.1587 | 17.1157 | 13.67% | 27.35% | 41.02% | 11.5304 | 12.7441
Zn, ppm | 95.7961 | 3.9131 | 87.9699 | 103.6224 | 84.0567 | 107.5355 | 4.085% | 8.17% | 12.25% | 91.0063 | 100.5859
Zr, ppm | 18.1371 | 3.0579 | 12.0213 | 24.2529 | 8.9634 | 27.3107 | 16.86% | 33.72% | 50.58% | 17.2302 | 19.0439
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 230b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PAAU02 | PhotonAssay (Au) two-cycle analysis
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 230b
Lab Slot | Laboratory Name and Location
Accura Geomysore Labs Private Limited, Jonnagiri, Kurnool, India
Accura Gold Minerals Testing, Namakkal, Tamilnadu, India
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Brisbane, QLD, Australia
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Britannia Mining Solutions, Hamilton, Ontario, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
Paragon Geochemical Laboratories, Surrey, BC, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Indo Mineral Research, Bungursari, West Java, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 230b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 2 | 1 | 290502 | 0.083723 | 0.38457 | 0.40043
2 | 1 | ANSLu | 1 | 18 | 5 | 290503 | 0.085852 | 0.35614 | 0.39891
2 | 1 | ANSLu | 1 | 6 | 2 | 290504 | 0.083976 | 0.39286 | 0.40087
2 | 1 | ANSLu | 1 | 10 | 3 | 290505 | 0.086992 | 0.4115 | 0.40186
2 | 1 | ANSLu | 1 | 7 | 8 | 290506 | 0.083164 | 0.3965 | 0.40106
2 | 1 | ANSLu | 1 | 15 | 10 | 290507 | 0.084909 | 0.40693 | 0.40162
2 | 1 | ANSLu | 1 | 11 | 9 | 290508 | 0.086816 | 0.41165 | 0.40187
2 | 1 | ANSLu | 1 | 13 | 4 | 290509 | 0.083285 | 0.41432 | 0.40201
2 | 1 | ANSLu | 1 | 4 | 7 | 290510 | 0.086639 | 0.39887 | 0.40119
2 | 1 | ANSLu | 1 | 17 | 5 | 290511 | 0.083773 | 0.37673 | 0.40001
2 | 1 | ANSLu | 1 | 14 | 4 | 290512 | 0.082623 | 0.42461 | 0.40256
2 | 1 | ANSLu | 1 | 16 | 10 | 290513 | 0.083712 | 0.36876 | 0.39958
2 | 1 | ANSLu | 1 | 8 | 8 | 290514 | 0.088098 | 0.402 | 0.40135
2 | 1 | ANSLu | 1 | 3 | 7 | 290515 | 0.085194 | 0.36946 | 0.39962
2 | 1 | ANSLu | 1 | 9 | 3 | 290516 | 0.085763 | 0.40181 | 0.40134
2 | 1 | ANSLu | 1 | 5 | 2 | 290517 | 0.084824 | 0.41704 | 0.40216
2 | 1 | ANSLu | 1 | 1 | 1 | 290518 | 0.085836 | 0.45356 | 0.4041
2 | 1 | ANSLu | 1 | 19 | 11 | 290519 | 0.086238 | 0.44236 | 0.40351
2 | 1 | ANSLu | 1 | 20 | 11 | 290520 | 0.08446 | 0.3971 | 0.40109
2 | 1 | ANSLu | 1 | 12 | 9 | 290521 | 0.087675 | 0.39961 | 0.40123
Mean | [6 empty] | 0.085178 | 0.40132 | 0.40132
Median | [7 empty] | 0.40071 | 0.40129
SD | [7 empty] | 0.023773 | 0.0012667
RSD | [7 empty] | 5.924% | 0.3156%
RSD Upscale Factor | [7 empty] | 0.053285
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 230b (Certified Value 0.381 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.30 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 25g | 30g | 50g | 30g | 30g | 30g | 50g | 40g | 30g | 30g | 30g | 30g | 30g | 50g | 40g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 3
| 1 | 1 | 0.45356 | 0.393 | 0.36594 | 0.374 | 0.388 | 0.38033 | 0.388 | 0.5 | 0.4 | 0.376 | 0.383 | 0.38 | 0.378 | 0.4128 | 0.35 | 0.39225 | 0.385 | 0.373 | 0.381 | 0.371 | 0.331 | 0.374 | 0.4 | 0.385 | 0.39 | 0.394 | 1
| 1 | 2 | 0.38457 | 0.393 | 0.36487 | 0.366 | 0.389 | 0.38347 | 0.37 | 0.46 | 0.4 | 0.372 | 0.382 | 0.36 | 0.374 | 0.4128 | 0.37 | 0.39799 | 0.374 | 0.373 | 0.384 | 0.371 | 0.335 | 0.369 | 0.394 | 0.379 | 0.37 | 0.4 | 0
| 1 | 3 | 0.36946 | 0.403 | 0.37343 | 0.371 | 0.378 | 0.38233 | 0.375 | 0.49 | 0.4 | 0.372 | 0.383 | 0.38 | 0.374 | 0.40248 | 0.37 | 0.394 | 0.384 | 0.372 | 0.401 | 0.367 | 0.323 | 0.391 | 0.394 | 0.38 | 0.38 | 0.403 | 16
| 1 | 4 | 0.39887 | 0.395 | 0.36808 | 0.371 | 0.361 | 0.38207 | 0.377 | 0.48 | 0.4 | 0.373 | 0.385 | 0.37 | 0.371 | 0.42168 | 0.36 | 0.39175 | 0.381 | 0.385 | 0.39 | 0.374 | 0.329 | 0.384 | 0.413 | 0.39 | 0.38 | 0.407 | 0.38099
| 1 | 5 | 0.41704 | 0.387 | 0.37343 | 0.378 | 0.367 | 0.38137 | 0.366 | 0.46 | 0.4 | 0.379 | 0.37 | 0.37 | 0.373 | 0.40248 | 0.37 | 0.38924 | 0.373 | 0.374 | 0.392 | 0.373 | 0.326 | 0.376 | 0.412 | 0.386 | 0.36 | 0.389 | 7
| 1 | 6 | 0.39286 | 0.396 | 0.37664 | 0.37 | 0.373 | 0.38223 | 0.387 | 0.48 | 0.4 | 0.367 | 0.384 | 0.37 | 0.375 | 0.42312 | 0.36 | 0.39449 | 0.374 | 0.375 | 0.389 | 0.377 | 0.33 | 0.381 | 0.4 | 0.39 | 0.36 | 0.396
|  | 7 | 2.00: 0.3965
|  | 8 | 2.00: 0.402
|  | 9 | 2.00: 0.40181
|  | 10 | 2.00: 0.4115
|  | 11 | 2.00: 0.41165
|  | 12 | 2.00: 0.39961
|  | 13 | 2.00: 0.41432
|  | 14 | 2.00: 0.42461
|  | 15 | 2.00: 0.40693
|  | 16 | 2.00: 0.36876
|  | 17 | 2.00: 0.37673
|  | 18 | 2.00: 0.35614
|  | 19 | 2.00: 0.44236
|  | 20 | 2.00: 0.3971
| Mean |  | 0.40132 | 0.3945 | 0.3704 | 0.37167 | 0.376 | 0.38197 | 0.37717 | 0.47833 | 0.4 | 0.37317 | 0.38117 | 0.37167 | 0.37417 | 0.41256 | 0.36333 | 0.39329 | 0.3785 | 0.37533 | 0.3895 | 0.37217 | 0.329 | 0.37917 | 0.40217 | 0.385 | 0.37333 | 0.39817
| Median |  | 0.40071 | 0.394 | 0.37076 | 0.371 | 0.3755 | 0.38215 | 0.376 | 0.48 | 0.4 | 0.3725 | 0.383 | 0.37 | 0.374 | 0.4128 | 0.365 | 0.39312 | 0.3775 | 0.3735 | 0.3895 | 0.372 | 0.3295 | 0.3785 | 0.4 | 0.3855 | 0.375 | 0.398
| Std Dev. |  | 0.023773 | 0.0052058 | 0.0047572 | 0.0040332 | 0.011243 | 0.0010511 | 0.0088863 | 0.016021 | 6.0809e-17 | 0.0040702 | 0.0055648 | 0.0075277 | 0.0023166 | 0.0089221 | 0.008165 | 0.0029622 | 0.0054681 | 0.0048442 | 0.0069498 | 0.0033714 | 0.0041473 | 0.0078337 | 0.0084479 | 0.0047329 | 0.012111 | 0.0064936
| Rel.Std.Dev. |  | 5.924% | 1.32% | 1.284% | 1.085% | 2.99% | 0.2752% | 2.356% | 3.349% | 1.52e-14% | 1.091% | 1.46% | 2.025% | 0.6191% | 2.163% | 2.247% | 0.7532% | 1.445% | 1.291% | 1.784% | 0.9059% | 1.261% | 2.066% | 2.101% | 1.229% | 3.244% | 1.631%
| PDM3 |  | 5.335% | 3.546% | -2.78% | -2.448% | -1.31% | 0.2556% | -1.004% | 25.55% | 4.989% | -2.054% | 0.04593% | -2.448% | -1.791% | 8.286% | -4.635% | 3.227% | -0.654% | -1.485% | 2.233% | -2.316% | -13.65% | -0.479% | 5.558% | 1.052% | -2.01% | 4.508%
| Z-Score (Absolute) |  | NA | 1.58 | 0.8 | 0.67 | 0.25 | 0.34 | 0.13 | 9.85 | 2.12 | 0.53 | 0.26 | 0.67 | 0.43 | 3.36 | 1.5 | 1.46 | 0 | 0.31 | 1.09 | 0.62 | 4.88 | 0.07 | 2.34 | 0.64 | 0.51 | 1.94
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 230b (Certified Value 0.374 ppm) | [14 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | ppm
|  |  | INAA | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | PAAU02 | 3
|  |  | 2.00: 0.085g | 3
| 1 | 1 | 0.45356 | 0.38 | 0.35 | 0.408 | 0.33 | 0.39 | 0.38 | 0.38 | 0.355 | 0.405 | 0.355 | 0.347 | 0.376 | 1
| 1 | 2 | 0.38457 | 0.37 | 0.37 | 0.415 | 0.35 | 0.38 | 0.36 | 0.396 | 0.349 | 0.383 | 0.36 | 0.371 | 0.364 | 0
| 1 | 3 | 0.36946 | 0.38 | 0.35 | 0.343 | 0.37 | 0.39 | 0.36 | 0.409 | 0.368 | 0.384 | 0.381 | 0.386 | 0.411 | 16
| 1 | 4 | 0.39887 | 0.4 | 0.36 | 0.386 | 0.34 | 0.36 | 0.38 | 0.354 | 0.388 | 0.341 | 0.37 | 0.364 | 0.401 | 0.37365
| 1 | 5 | 0.41704 | 0.4 | 0.38 | 0.331 | 0.35 | 0.38 | 0.4 | 0.388 | 0.374 | 0.403 | 0.37 | 0.356 | 0.366 | 9
| 1 | 6 | 0.39286 | 0.38 | 0.43 | 0.349 | 0.4 | 0.4 | 0.36 | 0.344 | 0.377 | 0.391 | 0.327 | 0.369 | 0.378
|  | 7 | 2.00: 0.3965
|  | 8 | 2.00: 0.402
|  | 9 | 2.00: 0.40181
|  | 10 | 2.00: 0.4115
|  | 11 | 2.00: 0.41165
|  | 12 | 2.00: 0.39961
|  | 13 | 2.00: 0.41432
|  | 14 | 2.00: 0.42461
|  | 15 | 2.00: 0.40693
|  | 16 | 2.00: 0.36876
|  | 17 | 2.00: 0.37673
|  | 18 | 2.00: 0.35614
|  | 19 | 2.00: 0.44236
|  | 20 | 2.00: 0.3971
| Mean |  | 0.40132 | 0.385 | 0.37333 | 0.372 | 0.35667 | 0.38333 | 0.37333 | 0.3785 | 0.3685 | 0.3845 | 0.3605 | 0.3655 | 0.38267
| Median |  | 0.40071 | 0.38 | 0.365 | 0.3675 | 0.35 | 0.385 | 0.37 | 0.384 | 0.371 | 0.3875 | 0.365 | 0.3665 | 0.377
| Std Dev. |  | 0.023773 | 0.012247 | 0.030111 | 0.035755 | 0.025033 | 0.013663 | 0.01633 | 0.024978 | 0.01446 | 0.023236 | 0.018727 | 0.013398 | 0.019138
| Rel.Std.Dev. |  | 5.924% | 3.181% | 8.065% | 9.611% | 7.019% | 3.564% | 4.374% | 6.599% | 3.924% | 6.043% | 5.195% | 3.666% | 5.001%
| PDM3 |  | 7.404% | 3.037% | -0.08549% | -0.4423% | -4.546% | 2.591% | -0.08549% | 1.297% | -1.379% | 2.903% | -3.52% | -2.182% | 2.412%
| Z-Score (Absolute) |  | NA | 0.92 | 0 | 0.1 | 1.31 | 0.79 | 0 | 0.41 | 0.38 | 0.88 | 1.01 | 0.62 | 0.73
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 230b (Certified Value 0.368 ppm) | [20 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.30 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | 3
|  |  | 0.085g | 25g | 30g | 30g | 50g | 30g | 25g | 25g | 25g | 25g | 25g | 30g | 50g | 25g | 30g | 25g | 10g | 15g | 30g | 3
| 1 | 1 | 0.45356 | 0.371 | 0.33793 | 0.382 | 0.44 | 0.39 | 0.351 | 0.379 | 0.37 | 0.351 | 0.352 | 0.42048 | 0.355 | 0.376 | 0.35 | 0.372 | 0.35 | 0.36 | 0.41 | 1
| 1 | 2 | 0.38457 | 0.359 | 0.33138 | 0.389 | 0.41 | 0.39 | 0.352 | 0.381 | 0.37 | 0.367 | 0.343 | 0.41676 | 0.35 | 0.35768 | 0.36 | 0.371 | 0.33 | 0.346 | 0.382 | 0
| 1 | 3 | 0.36946 | 0.372 | 0.33531 | 0.386 | 0.44 | 0.39 | 0.339 | 0.385 | 0.38 | 0.341 | 0.349 | 0.42336 | 0.36 | 0.38 | 0.36 | 0.365 | 0.31 | 0.359 | 0.413 | 16
| 1 | 4 | 0.39887 | 0.343 | 0.33531 | 0.388 | 0.44 | 0.39 | 0.351 | 0.367 | 0.38 | 0.342 | 0.357 | 0.40512 | 0.3625 | 0.374 | 0.35 | 0.37 | 0.35 | 0.351 | 0.399 | 0.36806
| 1 | 5 | 0.41704 | 0.377 | 0.33138 | 0.383 | 0.42 | 0.38 | 0.359 | 0.366 | 0.37 | 0.417 | 0.332 | 0.4176 | 0.355 | 0.38 | 0.35 | 0.369 | 0.36 | 0.357 | 0.386 | 11
| 1 | 6 | 0.39286 | 0.359 | 0.3471 | 0.391 | 0.48 | 0.38 | 0.358 | 0.368 | 0.37 | 0.335 | 0.338 | 0.41472 | 0.37 | 0.4088 | 0.36 | 0.364 | 0.35 | 0.361 | 0.404
|  | 7 | 2.00: 0.3965
|  | 8 | 2.00: 0.402
|  | 9 | 2.00: 0.40181
|  | 10 | 2.00: 0.4115
|  | 11 | 2.00: 0.41165
|  | 12 | 2.00: 0.39961
|  | 13 | 2.00: 0.41432
|  | 14 | 2.00: 0.42461
|  | 15 | 2.00: 0.40693
|  | 16 | 2.00: 0.36876
|  | 17 | 2.00: 0.37673
|  | 18 | 2.00: 0.35614
|  | 19 | 2.00: 0.44236
|  | 20 | 2.00: 0.3971
| Mean |  | 0.40132 | 0.3635 | 0.3364 | 0.3865 | 0.43833 | 0.38667 | 0.35167 | 0.37433 | 0.37333 | 0.35883 | 0.34517 | 0.41634 | 0.35875 | 0.37941 | 0.355 | 0.3685 | 0.34167 | 0.35567 | 0.399
| Median |  | 0.40071 | 0.365 | 0.33531 | 0.387 | 0.44 | 0.39 | 0.3515 | 0.3735 | 0.37 | 0.3465 | 0.346 | 0.41718 | 0.3575 | 0.378 | 0.355 | 0.3695 | 0.35 | 0.358 | 0.4015
| Std Dev. |  | 0.023773 | 0.012422 | 0.0058233 | 0.0035071 | 0.024014 | 0.005164 | 0.0071461 | 0.0082865 | 0.005164 | 0.030597 | 0.0092826 | 0.0062721 | 0.0070267 | 0.016599 | 0.0054772 | 0.0032711 | 0.018348 | 0.0059217 | 0.012649
| Rel.Std.Dev. |  | 5.924% | 3.417% | 1.731% | 0.9074% | 5.478% | 1.336% | 2.032% | 2.214% | 1.383% | 8.527% | 2.689% | 1.506% | 1.959% | 4.375% | 1.543% | 0.8877% | 5.37% | 1.665% | 3.17%
| PDM3 |  | 9.035% | -1.24% | -8.603% | 5.009% | 19.09% | 5.054% | -4.455% | 1.704% | 1.432% | -2.508% | -6.221% | 13.12% | -2.53% | 3.084% | -3.549% | 0.1187% | -7.172% | -3.368% | 8.405%
| Z-Score (Absolute) |  | NA | 0.12 | 1.44 | 1 | 3.52 | 1 | 0.7 | 0.41 | 0.36 | 0.35 | 1.01 | 2.45 | 0.35 | 0.65 | 0.53 | 0.12 | 1.18 | 0.5 | 1.6
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 230b (Certified Value 0.363 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*MS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 200g | 50g | 200g | 50g | 30g | 30g | 30g | 100g | 50g | 10g | 05g | 05g | 15g | 50g | 15g | 100g | 3
| 1 | 1 | 0.45356 | 0.36 | 0.37448 | 0.2904 | 0.315 | 0.33 | 0.37924 | 0.323 | 0.36 | 0.37 | 0.4 | 0.36 | 0.32 | 0.39456 | 0.35 | < 0.01 | 0.3 | 0.366 | 0.37 | 0.241 | 0.39 | 0.405 | 1
| 1 | 2 | 0.38457 | 0.36 | 0.3701 | 0.2944 | 0.34 | 0.34 | 0.37277 | 0.311 | 0.35 | 0.36 | 0.39 | 0.37 | 0.3 | 0.39648 | 0.34598 | < 0.01 | 0.3 | 0.363 | 0.36 | 0.251 | 0.37 | 0.407 | 0
| 1 | 3 | 0.36946 | 0.36 | 0.37726 | 0.2976 | 0.341 | 0.321 | 0.37263 | 0.321 | 0.36 | 0.36 | 0.39 | 0.37 | 0.42 | 0.40512 | 0.34978 | < 0.01 | 0.3 | 0.358 | 0.38 | 0.233 | 0.37 | 0.403 | 16
| 1 | 4 | 0.39887 | 0.36 | 0.37036 | 0.2608 | 0.572 | 0.34 | 0.36813 | 0.298 | 0.34 | 0.37 | 0.38 | 0.38 | 0.32 | 0.40512 | 0.36874 | < 0.01 | 0.3 | 0.362 | 0.35 | 0.233 | 0.38 | 0.402 | 0.36314
| 1 | 5 | 0.41704 | 0.36 | 0.36769 | 0.288 | 0.32 | 0.33 | 0.36789 | 0.352 | 0.37 | 0.37 | 0.39 | 0.31 | 0.39 | 0.40608 | 0.35 | < 0.01 | 0.3 | 0.361 | 0.36 | 0.227 | 0.37 | 0.402 | 13
| 1 | 6 | 0.39286 | 0.35 | 0.37478 | 0.28 | 0.345 | 0.33 | 0.38221 | 0.326 | 0.35 | 0.37 | 0.39 | 0.34 | 0.39 | 0.39552 | 0.36793 | < 0.01 | 0.3 | 0.365 | 0.36 | 0.236 | 0.36 | 0.408
|  | 7 | 2.00: 0.3965
|  | 8 | 2.00: 0.402
|  | 9 | 2.00: 0.40181
|  | 10 | 2.00: 0.4115
|  | 11 | 2.00: 0.41165
|  | 12 | 2.00: 0.39961
|  | 13 | 2.00: 0.41432
|  | 14 | 2.00: 0.42461
|  | 15 | 2.00: 0.40693
|  | 16 | 2.00: 0.36876
|  | 17 | 2.00: 0.37673
|  | 18 | 2.00: 0.35614
|  | 19 | 2.00: 0.44236
|  | 20 | 2.00: 0.3971
| Mean |  | 0.40132 | 0.35833 | 0.37244 | 0.2852 | 0.37217 | 0.33183 | 0.37381 | 0.32183 | 0.355 | 0.36667 | 0.39 | 0.355 | 0.35667 | 0.40048 | 0.3554 |  | 0.3 | 0.3625 | 0.36333 | 0.23683 | 0.37333 | 0.4045
| Median |  | 0.40071 | 0.36 | 0.37242 | 0.2892 | 0.3405 | 0.33 | 0.3727 | 0.322 | 0.355 | 0.37 | 0.39 | 0.365 | 0.355 | 0.4008 | 0.35 |  | 0.3 | 0.3625 | 0.36 | 0.2345 | 0.37 | 0.404
| Std Dev. |  | 0.023773 | 0.0040825 | 0.0036107 | 0.013384 | 0.098658 | 0.0072226 | 0.0058299 | 0.017971 | 0.010488 | 0.005164 | 0.0063246 | 0.025884 | 0.049261 | 0.0054785 | 0.010134 |  | 0 | 0.002881 | 0.010328 | 0.0083046 | 0.010328 | 0.0025884
| Rel.Std.Dev. |  | 5.924% | 1.139% | 0.9695% | 4.693% | 26.51% | 2.177% | 1.56% | 5.584% | 2.954% | 1.408% | 1.622% | 7.291% | 13.81% | 1.368% | 2.851% |  | 0% | 0.7948% | 2.843% | 3.507% | 2.766% | 0.6399%
| PDM3 |  | 10.51% | -1.323% | 2.563% | -21.46% | 2.486% | -8.62% | 2.939% | -11.37% | -2.241% | 0.9718% | 7.397% | -2.241% | -1.782% | 10.28% | -2.13% |  | -17.39% | -0.1756% | 0.05391% | -34.78% | 2.808% | 11.39%
| Z-Score (Absolute) |  | NA | 0.28 | 0.67 | 5.24 | 0.66 | 2.08 | 0.77 | 2.76 | 0.51 | 0.28 | 1.86 | 0.51 | 0.4 | 2.58 | 0.48 | NA | NA | 0 | 0.06 | 8.52 | 0.73 | 2.85
Results from laboratory 18 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 230b (Certified Value 0.195 ppm) | [18 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.21 | < 1 | 0.21677 | < 0.3 | 0.31 | 0.19 | 0.19 | 0.39 | 0.2 | 0.16 | 0.61983 | 0.2 | 0.19 | 0.19 | < 0.3 | 1.11 | 0.2 | 1
| 1 | 2 | 0.18 | < 1 | 0.20335 | < 0.3 | 0.32 | 0.19 | 0.2 | 0.22 | 0.2 | 0.17 | 0.59882 | 0.2 | 0.21 | 0.18 | < 0.3 | 1.07 | 0.2 | 19
| 1 | 3 | 0.19 | < 1 | 0.20812 | < 0.3 | 0.34 | 0.2 | 0.19 | 0.21 | 0.19 | 0.18 | 0.5943 | 0.2 | 0.18 | 0.18 | < 0.3 | 1.22 | 0.2 | 16
| 1 | 4 | 0.19 | < 1 | 0.21386 | < 0.3 | 0.33 | 0.18 | 0.18 | 0.22 | 0.2 | 0.18 | 0.54634 | 0.2 | 0.21 | 0.19 | < 0.3 | 1.15 | 0.2 | 0.19483
| 1 | 5 | 0.17 | < 1 | 0.23217 | < 0.3 | 0.34 | 0.2 | 0.17 | 0.21 | 0.21 | 0.17 | 0.5942 | 0.2 | 0.2 | 0.18 | < 0.3 | 1.02 | 0.2 | 15
| 1 | 6 | 0.18 | < 1 | 0.2188 | < 0.3 | 0.34 | 0.19 | 0.2 | 0.23 | 0.19 | 0.17 | 0.52668 | 0.2 | 0.21 | 0.18 | < 0.3 | 1.03 | 0.2
| Mean |  | 0.18667 |  | 0.21551 |  | 0.33 | 0.19167 | 0.18833 | 0.24667 | 0.19833 | 0.17167 | 0.58003 | 0.2 | 0.2 | 0.18333 |  | 1.1 | 0.2
| Median |  | 0.185 |  | 0.21532 |  | 0.335 | 0.19 | 0.19 | 0.22 | 0.2 | 0.17 | 0.59425 | 0.2 | 0.205 | 0.18 |  | 1.09 | 0.2
| Std Dev. |  | 0.013663 |  | 0.0099512 |  | 0.012649 | 0.0075277 | 0.01169 | 0.070616 | 0.0075277 | 0.0075277 | 0.035561 | 3.0405e-17 | 0.012649 | 0.005164 |  | 0.07642 | 3.0405e-17
| Rel.Std.Dev. |  | 7.319% |  | 4.617% |  | 3.833% | 3.928% | 6.207% | 28.63% | 3.795% | 4.385% | 6.131% | 1.52e-14% | 6.325% | 2.817% |  | 6.947% | 1.52e-14%
| PDM3 |  | -4.192% |  | 10.61% |  | 69.37% | -1.626% | -3.337% | 26.6% | 1.796% | -11.89% | 197.7% | 2.651% | 2.651% | -5.903% |  | 464.6% | 2.651%
| Z-Score (Absolute) |  | 0.3 | 7.63 | 0.43 | 1.22 | 3.33 | 0.17 | 0.25 | 1.22 | 0 | 0.67 | 9.65 | NA | 0.04 | 0.38 | 1.22 | 22.8 | NA
Results from laboratories 18 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 230b (Certified Value 6.82 wt.%) | [18 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 6.9644 | 7.0455 | 6.899 | 5.85 | 6.77 | 6.91 | 6.86 | 6.61 | 6.79 | 6.71 | 9.2235 | 6.74 | 6.65 | 6.9539 | 6.1348 | 6.98 | 6.84 | 1
| 1 | 2 | 7.011 | 7.0334 | 6.9171 | 6.23 | 6.64 | 6.84 | 6.91 | 6.5 | 6.78 | 6.81 | 9.0913 | 6.72 | 6.71 | 7.0661 | 6.2873 | 6.91 | 6.76 | 0
| 1 | 3 | 6.9975 | 7.0453 | 7.1134 | 6.77 | 6.67 | 6.84 | 6.88 | 6.66 | 6.75 | 6.85 | 9.1314 | 6.76 | 6.66 | 7.031 | 6.2642 | 6.97 | 6.81 | 16
| 1 | 4 | 6.953 | 7.0232 | 6.8938 | 6.47 | 6.62 | 6.75 | 6.98 | 6.52 | 6.64 | 6.82 | 9.2074 | 6.78 | 6.75 | 7.1064 | 6.29 | 7.07 | 6.76 | 6.8164
| 1 | 5 | 6.9232 | 7.0614 | 6.9009 | 6.86 | 6.45 | 6.71 | 6.77 | 6.56 | 6.67 | 6.85 | 9.6294 | 6.81 | 6.77 | 7.0705 | 6.1724 | 7.08 | 6.73 | 16
| 1 | 6 | 7.0645 | 7.0577 | 6.8734 | 6.33 | 6.76 | 6.71 | 7.16 | 6.56 | 6.6 | 6.84 | 9.193 | 6.78 | 6.72 | 7.175 | 6.1583 | 6.94 | 6.65
| Mean |  | 6.9856 | 7.0444 | 6.9329 | 6.4183 | 6.6517 | 6.7933 | 6.9267 | 6.5683 | 6.705 | 6.8133 | 9.246 | 6.765 | 6.71 | 7.0671 | 6.2178 | 6.9917 | 6.7583
| Median |  | 6.9809 | 7.0454 | 6.9 | 6.4 | 6.655 | 6.795 | 6.895 | 6.56 | 6.71 | 6.83 | 9.2002 | 6.77 | 6.715 | 7.0683 | 6.2183 | 6.975 | 6.76
| Std Dev. |  | 0.049836 | 0.01441 | 0.089506 | 0.37086 | 0.11652 | 0.082138 | 0.13322 | 0.058793 | 0.079183 | 0.053166 | 0.19433 | 0.032094 | 0.047749 | 0.073915 | 0.070266 | 0.069113 | 0.066156
| Rel.Std.Dev. |  | 0.7134% | 0.2046% | 1.291% | 5.778% | 1.752% | 1.209% | 1.923% | 0.8951% | 1.181% | 0.7803% | 2.102% | 0.4744% | 0.7116% | 1.046% | 1.13% | 0.9885% | 0.9789%
| PDM3 |  | 2.483% | 3.346% | 1.71% | -5.839% | -2.416% | -0.3378% | 1.618% | -3.639% | -1.634% | -0.04444% | 35.64% | -0.7535% | -1.56% | 3.679% | -8.781% | 2.572% | -0.8513%
| Z-Score (Absolute) |  | 0.92 | 1.2 | 0.66 | 1.78 | 0.67 | 0 | 0.63 | 1.07 | 0.42 | 0.1 | 11.67 | 0.13 | 0.4 | 1.3 | 2.74 | 0.94 | 0.17
| Analytical results for As in OREAS 230b (Certified Value 26.8 ppm) | [18 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 1
|  |  |  | 0.25g | [15 empty] | 2
| 1 | 1 | 26.6 | 26.456 | 27.2887 | 27.7 | 17 | 26.4 | 28.5 | 26.9 | 27.4 | 25 | 28.7858 | 25 | 24 | 26.4 | 25.7 | 1.4 | 12.1 | 1
| 1 | 2 | 26.2 | 25.862 | 26.2469 | 28.1 | 17.8 | 26.1 | 28.8 | 28.1 | 27.3 | 25.7 | 29.1958 | 26 | 24 | 27 | 26.3 | 2.65 | 12.2 | 20
| 1 | 3 | 26.7 | 26.775 | 27.3842 | 28.2 | 17.9 | 25.9 | 29.2 | 27 | 26.1 | 25.9 | 30.0572 | 25 | 25 | 26.7 | 26.6 | 2.44 | 12 | 16
| 1 | 4 | 26.4 | 27.93 | 26.9142 | 28.1 | 17 | 25.6 | 28.6 | 27.3 | 27.3 | 25.9 | 29.6543 | 26 | 25 | 27.2 | 25.7 | 2.28 | 13 | 26.8227
| 1 | 5 | 26.8 | 28.612 | 26.5338 | 27.6 | 17.6 | 25.6 | 27.8 | 27.6 | 28.7 | 25.4 | 27.59 | 26 | 26 | 26.4 | 26.3 | 2.43 | 12.4 | 17
| 1 | 6 | 26.3 | 28.755 | 26.5364 | 28.7 | 17.5 | 25.6 | 30.6 | 27.4 | 26.1 | 26.2 | 28.4512 | 25 | 25 | 26.7 | 25.7 | 1.41 | 14.7
| Mean |  | 26.5 | 27.3983 | 26.8174 | 28.0667 | 17.4667 | 25.8667 | 28.9167 | 27.3833 | 27.15 | 25.6833 | 28.9557 | 25.5 | 24.8333 | 26.7333 | 26.05 | 2.1017 | 12.7333
| Median |  | 26.5 | 27.3525 | 26.7253 | 28.1 | 17.55 | 25.75 | 28.7 | 27.35 | 27.3 | 25.8 | 28.9908 | 25.5 | 25 | 26.7 | 26 | 2.355 | 12.3
| Std Dev. |  | 0.23664 | 1.2028 | 0.45552 | 0.39328 | 0.38816 | 0.33267 | 0.94322 | 0.43551 | 0.97108 | 0.42622 | 0.88405 | 0.54772 | 0.75277 | 0.32042 | 0.39875 | 0.55236 | 1.027
| Rel.Std.Dev. |  | 0.893% | 4.39% | 1.699% | 1.401% | 2.222% | 1.286% | 3.262% | 1.59% | 3.577% | 1.66% | 3.053% | 2.148% | 3.031% | 1.199% | 1.531% | 26.28% | 8.065%
| PDM3 |  | -1.203% | 2.146% | -0.02001% | 4.638% | -34.88% | -3.564% | 7.807% | 2.09% | 1.22% | -4.248% | 7.952% | -4.931% | -7.417% | -0.3332% | -2.881% | -92.16% | -52.53%
| Z-Score (Absolute) |  | 0.09 | 0.61 | 0.16 | 1.14 | 7.2 | 0.59 | 1.81 | 0.6 | 0.42 | 0.73 | 1.84 | 0.88 | 1.4 | 0.09 | 0.45 | NA | 10.92
| Analytical results for Ba in OREAS 230b (Certified Value 94 ppm) | [18 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 100.7 | 90.71 | 93.8589 | 92 | 98 | 100 | 100 | 90 | 100 | 90 | 95.4192 | 90 | 96 | 96.2 | 84.61 | 122 |  | 1
| 1 | 2 | 97.5 | 89.43 | 92.5868 | 94 | 93 | 100 | 100 | 90 | 100 | 100 | 95.2352 | 91 | 98 | 95.7 | 86.49 | 128 |  | 21
| 1 | 3 | 97.2 | 87.97 | 93.4749 | 118 | 98 | 100 | 100 | 90 | 100 | 90 | 98.3773 | 89 | 95 | 95.5 | 85.95 | 124 |  | 16
| 1 | 4 | 97.2 | 88.43 | 94.357 | 92 | 96 | 90 | 100 | 90 | 90 | 90 | 96.6893 | 90 | 96 | 97 | 86.24 | 126 |  | 94.1885
| 1 | 5 | 96.1 | 88.55 | 93.2401 | 93 | 97 | 90 | 100 | 90 | 90 | 90 | 91.5648 | 91 | 97 | 96.6 | 85.41 | 122 |  | 18
| 1 | 6 | 95.7 | 88.7 | 92.9025 | 92 | 100 | 90 | 100 | 90 | 90 | 90 | 92.4016 | 90 | 100 | 95.9 | 87.08 | 124
| Mean |  | 97.4 | 88.965 | 93.4033 | 96.8333 | 97 | 95 | 100 | 90 | 95 | 91.6667 | 94.9479 | 90.1667 | 97 | 96.15 | 85.9633 | 124.3333
| Median |  | 97.2 | 88.625 | 93.3575 | 92.5 | 97.5 | 95 | 100 | 90 | 95 | 90 | 95.3272 | 90 | 96.5 | 96.05 | 86.095 | 124
| Std Dev. |  | 1.7641 | 0.97775 | 0.64286 | 10.4003 | 2.3664 | 5.4772 | 0 | 0 | 5.4772 | 4.0825 | 2.5701 | 0.75277 | 1.7889 | 0.56833 | 0.86472 | 2.3381
| Rel.Std.Dev. |  | 1.811% | 1.099% | 0.6883% | 10.74% | 2.44% | 5.766% | 0% | 0% | 5.766% | 4.454% | 2.707% | 0.8349% | 1.844% | 0.5911% | 1.006% | 1.881%
| PDM3 |  | 3.41% | -5.546% | -0.8336% | 2.808% | 2.985% | 0.8615% | 6.17% | -4.447% | 0.8615% | -2.677% | 0.8062% | -4.27% | 2.985% | 2.082% | -8.733% | 32%
| Z-Score (Absolute) |  | 0.74 | 1.85 | 0.49 | 0.56 | 0.61 | 0 | 1.53 | 1.53 | 0 | 1.02 | 0.02 | 1.48 | 0.61 | 0.35 | 2.77 | 8.99
| Analytical results for Be in OREAS 230b (Certified Value 0.37 ppm) | [18 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.41 | 0.31712 | 0.40454 | 0.5 | 0.42 | 0.36 | 0.37 | 0.34 | 0.37 | 0.37 | 0.37782 | < 0.5 | 0.4 | 0.33 | 0.34 | < 1 |  | 1
| 1 | 2 | 0.41 | 0.36376 | 0.39501 | 0.5 | 0.42 | 0.36 | 0.37 | 0.32 | 0.36 | 0.37 | 0.3972 | < 0.5 | 0.5 | 0.36 | 0.37 | < 1 |  | 22
| 1 | 3 | 0.37 | 0.36542 | 0.4052 | 0.4 | 0.4 | 0.37 | 0.37 | 0.35 | 0.36 | 0.38 | 0.38721 | < 0.5 | 0.4 | 0.38 | 0.35 | < 1 |  | 16
| 1 | 4 | 0.37 | 0.35345 | 0.39283 | 0.4 | 0.39 | 0.36 | 0.38 | 0.32 | 0.36 | 0.38 | 0.38481 | < 0.5 | 0.4 | 0.38 | 0.34 | < 1 |  | 0.37266
| 1 | 5 | 0.42 | 0.34759 | 0.39993 | 0.5 | 0.4 | 0.35 | 0.37 | 0.33 | 0.36 | 0.38 | 0.38162 | < 0.5 | 0.5 | 0.37 | 0.35 | < 1 |  | 19
| 1 | 6 | 0.38 | 0.36379 | 0.40251 | 0.6 | 0.42 | 0.36 | 0.38 | 0.33 | 0.36 | 0.38 | 0.35124 | < 0.5 | 0.5 | 0.38 | 0.34 | < 1
| Mean |  | 0.39333 | 0.35186 | 0.4 | 0.48333 | 0.40833 | 0.36 | 0.37333 | 0.33167 | 0.36167 | 0.37667 | 0.37998 |  | 0.45 | 0.36667 | 0.34833
| Median |  | 0.395 | 0.35861 | 0.40122 | 0.5 | 0.41 | 0.36 | 0.37 | 0.33 | 0.36 | 0.38 | 0.38322 |  | 0.45 | 0.375 | 0.345
| Std Dev. |  | 0.022509 | 0.018415 | 0.0051028 | 0.075277 | 0.013292 | 0.0063246 | 0.005164 | 0.01169 | 0.0040825 | 0.005164 | 0.015527 |  | 0.054772 | 0.019664 | 0.01169
| Rel.Std.Dev. |  | 5.723% | 5.234% | 1.276% | 15.57% | 3.255% | 1.757% | 1.383% | 3.525% | 1.129% | 1.371% | 4.086% |  | 12.17% | 5.363% | 3.356%
| PDM3 |  | 5.549% | -5.582% | 7.339% | 29.7% | 9.574% | -3.396% | 0.1818% | -11% | -2.949% | 1.076% | 1.966% |  | 20.75% | -1.607% | -6.527%
| Z-Score (Absolute) |  | 0.79 | 0.61 | 1.02 | NA | 1.3 | 0.34 | 0.11 | 1.3 | 0.28 | 0.23 | 0.34 | 4.07 | NA | 0.11 | 0.73 | 4.4
Results from laboratories 7 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 230b (Certified Value 0.23 ppm) | [18 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  |  | 0.25g | [15 empty] | 3
| 1 | 1 | 0.25 | 0.1691 | 0.21951 | 0.23 | 1.21 | 0.22 | 0.21 | 0.23 | 0.26 | 0.22 | 0.22175 | 0.22 | 0.26 | 0.25 | 0.2 | < 0.5 | 0.2 | 1
| 1 | 2 | 0.24 | 0.16102 | 0.22255 | 0.25 | 1.12 | 0.21 | 0.23 | 0.24 | 0.25 | 0.23 | 0.21805 | 0.22 | 0.27 | 0.24 | 0.16 | < 0.5 | 0.2 | 23
| 1 | 3 | 0.25 | 0.17607 | 0.23317 | 0.24 | 1.12 | 0.22 | 0.22 | 0.24 | 0.24 | 0.23 | 0.22948 | 0.22 | 0.26 | 0.24 | 0.13 | < 0.5 | 0.3 | 16
| 1 | 4 | 0.25 | 0.14916 | 0.212 | 0.23 | 1.23 | 0.22 | 0.22 | 0.23 | 0.26 | 0.23 | 0.22269 | 0.23 | 0.29 | 0.25 | 0.21 | < 0.5 | 0.3 | 0.23144
| 1 | 5 | 0.25 | 0.16146 | 0.21261 | 0.23 | 1.16 | 0.22 | 0.2 | 0.24 | 0.25 | 0.23 | 0.20046 | 0.22 | 0.26 | 0.25 | 0.28 | < 0.5 | 0.2 | 20
| 1 | 6 | 0.25 | 0.17751 | 0.22098 | 0.24 | 1.19 | 0.22 | 0.21 | 0.23 | 0.24 | 0.22 | 0.20385 | 0.22 | 0.25 | 0.26 | 0.33 | < 0.5 | 0.3
| Mean |  | 0.24833 | 0.16572 | 0.22013 | 0.23667 | 1.1717 | 0.21833 | 0.215 | 0.235 | 0.25 | 0.22667 | 0.21605 | 0.22167 | 0.265 | 0.24833 | 0.21833 |  | 0.25
| Median |  | 0.25 | 0.16528 | 0.22024 | 0.235 | 1.175 | 0.22 | 0.215 | 0.235 | 0.25 | 0.23 | 0.2199 | 0.22 | 0.26 | 0.25 | 0.205 |  | 0.25
| Std Dev. |  | 0.0040825 | 0.010696 | 0.0077402 | 0.008165 | 0.046224 | 0.0040825 | 0.010488 | 0.0054772 | 0.0089443 | 0.005164 | 0.011427 | 0.0040825 | 0.013784 | 0.0075277 | 0.074677 |  | 0.054772
| Rel.Std.Dev. |  | 1.644% | 6.454% | 3.516% | 3.45% | 3.945% | 1.87% | 4.878% | 2.331% | 3.578% | 2.278% | 5.289% | 1.842% | 5.202% | 3.031% | 34.2% |  | 21.91%
| PDM3 |  | 7.301% | -28.4% | -4.883% | 2.26% | 406.3% | -5.662% | -7.102% | 1.54% | 8.021% | -2.061% | -6.65% | -4.222% | 14.5% | 7.301% | -5.662% |  | 8.021%
| Z-Score (Absolute) |  | 0.77 | 2.87 | 0.47 | 0.26 | 41.44 | 0.55 | 0.7 | 0.18 | 0.84 | 0.18 | 0.65 | 0.4 | 1.5 | 0.77 | 0.55 | 0.84 | NA
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 230b (Certified Value 7.22 wt.%) | [18 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 7.2837 | 7.2696 | 7.2229 | 7.42 | 7.27 | 7.3 | 7.37 | 6.96 | 7.22 | 6.97 | 9.4764 | 7.23 | 6.83 | 7.1178 | 7.1749 | 6.72 | 7.55 | 1
| 1 | 2 | 7.3807 | 7.2741 | 7.1299 | 7.52 | 7.15 | 7.36 | 7.42 | 6.87 | 7.14 | 7.27 | 9.7185 | 7.27 | 6.8 | 7.2443 | 7.245 | 6.76 | 7.57 | 0
| 1 | 3 | 7.3715 | 7.259 | 7.3141 | 7.36 | 7.19 | 7.44 | 7.43 | 7.09 | 7.08 | 7.11 | 9.9857 | 7.23 | 6.86 | 7.1978 | 7.1919 | 6.69 | 7.62 | 16
| 1 | 4 | 7.2806 | 7.2506 | 7.1042 | 7.44 | 7.15 | 7.19 | 7.53 | 6.89 | 6.97 | 7.1 | 9.9939 | 7.25 | 6.91 | 7.3228 | 7.2457 | 6.8 | 7.62 | 7.2202
| 1 | 5 | 7.2242 | 7.2563 | 7.1662 | 7.28 | 6.96 | 7.07 | 7.33 | 6.91 | 6.98 | 7.21 | 9.3954 | 7.29 | 6.81 | 7.2492 | 7.1347 | 6.82 | 7.62 | 21
| 1 | 6 | 7.391 | 7.2691 | 7.1647 | 7.45 | 7.27 | 7.23 | 7.71 | 6.94 | 6.97 | 7.14 | 9.5608 | 7.25 | 6.87 | 7.3271 | 7.2225 | 6.8 | 7.6
| Mean |  | 7.322 | 7.2631 | 7.1837 | 7.4117 | 7.165 | 7.265 | 7.465 | 6.9433 | 7.06 | 7.1333 | 9.6885 | 7.2533 | 6.8467 | 7.2432 | 7.2024 | 6.765 | 7.5967
| Median |  | 7.3276 | 7.264 | 7.1655 | 7.43 | 7.17 | 7.265 | 7.425 | 6.925 | 7.03 | 7.125 | 9.6396 | 7.25 | 6.845 | 7.2467 | 7.2072 | 6.78 | 7.61
| Std Dev. |  | 0.068414 | 0.0091623 | 0.075391 | 0.082563 | 0.11415 | 0.13096 | 0.13766 | 0.078909 | 0.10488 | 0.10289 | 0.25681 | 0.023381 | 0.041312 | 0.078992 | 0.043689 | 0.051284 | 0.030111
| Rel.Std.Dev. |  | 0.9344% | 0.1261% | 1.049% | 1.114% | 1.593% | 1.803% | 1.844% | 1.136% | 1.486% | 1.442% | 2.651% | 0.3223% | 0.6034% | 1.091% | 0.6066% | 0.7581% | 0.3964%
| PDM3 |  | 1.41% | 0.5949% | -0.5054% | 2.652% | -0.7639% | 0.6211% | 3.391% | -3.834% | -2.218% | -1.202% | 34.19% | 0.4595% | -5.173% | 0.3187% | -0.2452% | -6.304% | 5.215%
| Z-Score (Absolute) |  | 0.71 | 0.29 | 0.28 | 1.35 | 0.41 | 0.3 | 1.73 | 2 | 1.16 | 0.64 | NA | 0.22 | 2.69 | 0.15 | 0.15 | 3.27 | 2.67
| Analytical results for Cd in OREAS 230b (Certified Value 0.31 ppm) | [18 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 0.31 | 3.3 | 0.26176 | < 0.3 | 0.29 | 0.31 | 0.32 | 0.31 | 0.31 | 0.3 | 0.73367 | 0.3 | 0.26 | 0.29 | 0.5 | 0.31 | 0.3 | 1
| 1 | 2 | 0.3 | 3.32 | 0.3025 | < 0.3 | 0.28 | 0.29 | 0.31 | 0.29 | 0.31 | 0.3 | 0.64309 | 0.29 | 0.27 | 0.33 | 0.13 | 0.3 | 0.3 | 24
| 1 | 3 | 0.29 | 3.27 | 0.29047 | < 0.3 | 0.29 | 0.3 | 0.31 | 0.39 | 0.32 | 0.28 | 0.75432 | 0.31 | 0.3 | 0.33 | 0.38 | 0.3 | 0.3 | 16
| 1 | 4 | 0.3 | 3.56 | 0.25598 | < 0.3 | 0.25 | 0.29 | 0.3 | 0.34 | 0.34 | 0.3 | 0.69664 | 0.31 | 0.31 | 0.32 | 0.38 | 0.32 | 0.4 | 0.30789
| 1 | 5 | 0.3 | 3.46 | 0.28559 | < 0.3 | 0.28 | 0.29 | 0.29 | 0.34 | 0.32 | 0.32 | 0.66625 | 0.31 | 0.29 | 0.26 | 1.01 | 0.32 | 0.4 | 22
| 1 | 6 | 0.28 | 3.12 | 0.29947 | < 0.3 | 0.3 | 0.3 | 0.32 | 0.32 | 0.3 | 0.31 | 0.65601 | 0.3 | 0.28 | 0.29 | < 0.02 | 0.34 | 0.3
| Mean |  | 0.29667 | 3.3383 | 0.28263 |  | 0.28167 | 0.29667 | 0.30833 | 0.33167 | 0.31667 | 0.30167 | 0.69166 | 0.30333 | 0.285 | 0.30333 | 0.48 | 0.315 | 0.33333
| Median |  | 0.3 | 3.31 | 0.28803 |  | 0.285 | 0.295 | 0.31 | 0.33 | 0.315 | 0.3 | 0.68145 | 0.305 | 0.285 | 0.305 | 0.38 | 0.315 | 0.3
| Std Dev. |  | 0.010328 | 0.15368 | 0.019465 |  | 0.017224 | 0.008165 | 0.01169 | 0.034303 | 0.013663 | 0.013292 | 0.044702 | 0.008165 | 0.018708 | 0.028048 | 0.3255 | 0.015166 | 0.05164
| Rel.Std.Dev. |  | 3.481% | 4.603% | 6.887% |  | 6.115% | 2.752% | 3.791% | 10.34% | 4.315% | 4.406% | 6.463% | 2.692% | 6.564% | 9.246% | 67.81% | 4.815% | 15.49%
| PDM3 |  | -3.647% | 984.2% | -8.206% |  | -8.518% | -3.647% | 0.1426% | 7.721% | 2.849% | -2.023% | 124.6% | -1.481% | -7.436% | -1.481% | 55.9% | 2.308% | 8.262%
| Z-Score (Absolute) |  | 0.34 | NA | 1.05 | 7.75 | 1.1 | 0.34 | 0.25 | 1.43 | 0.67 | 0.08 | 19.64 | 0 | 0.93 | 0 | 4.97 | 0.59 | NA
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 230b (Certified Value 11.1 ppm) | [18 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 10.86 | 10.9188 | 10.7721 | 10.8 | 10.9 | 11.45 | 12 | 10.3 | 12.05 | 12 | 13.4129 | 11.39 | 10.23 | 8.94 | 19 |  |  | 1
| 1 | 2 | 10.79 | 10.8342 | 10.8642 | 10.5 | 10.69 | 11.45 | 11.95 | 9.97 | 11.9 | 11 | 13.3548 | 11.85 | 10.49 | 8.16 | 15 |  |  | 25
| 1 | 3 | 10.93 | 10.7107 | 10.7383 | 10.5 | 10.54 | 11.7 | 12.2 | 10.15 | 11.95 | 11.5 | 13.7803 | 11.62 | 10.14 | 8.3 | 19 |  |  | 16
| 1 | 4 | 10.87 | 10.8352 | 10.3045 | 10.7 | 10.82 | 11.55 | 11.55 | 10.1 | 12.25 | 12.3 | 13.4497 | 11.35 | 10.33 | 8.29 | 19 |  |  | 11.06
| 1 | 5 | 11.25 | 10.71 | 10.3658 | 10.3 | 10.96 | 11.45 | 11.4 | 10.05 | 12.2 | 11.4 | 12.8291 | 11.58 | 10.2 | 8.76 | 15 |  |  | 23
| 1 | 6 | 10.87 | 10.2259 | 10.3278 | 10.8 | 11.03 | 11.35 | 12.25 | 10.05 | 11.7 | 12.3 | 13.001 | 11.41 | 10.54 | 8.51 | 14
| Mean |  | 10.9283 | 10.7058 | 10.5621 | 10.6 | 10.8233 | 11.4917 | 11.8917 | 10.1033 | 12.0083 | 11.75 | 13.3046 | 11.5333 | 10.3217 | 8.4933 | 16.8333
| Median |  | 10.87 | 10.7725 | 10.5521 | 10.6 | 10.86 | 11.45 | 11.975 | 10.075 | 12 | 11.75 | 13.3838 | 11.495 | 10.28 | 8.405 | 17
| Std Dev. |  | 0.16376 | 0.24858 | 0.25544 | 0.2 | 0.18184 | 0.12007 | 0.34557 | 0.11343 | 0.20351 | 0.53198 | 0.34053 | 0.18938 | 0.16266 | 0.30356 | 2.4014
| Rel.Std.Dev. |  | 1.498% | 2.322% | 2.418% | 1.887% | 1.68% | 1.045% | 2.906% | 1.123% | 1.695% | 4.527% | 2.56% | 1.642% | 1.576% | 3.574% | 14.27%
| PDM3 |  | -1.19% | -3.202% | -4.501% | -4.159% | -2.14% | 3.903% | 7.52% | -8.65% | 8.575% | 6.239% | 20.3% | 4.28% | -6.675% | -23.21% | 52.2%
| Z-Score (Absolute) |  | 0.06 | 0.18 | 0.33 | 0.29 | 0.06 | 0.65 | 1.08 | 0.82 | 1.2 | 0.93 | 2.57 | 0.7 | 0.59 | 2.52 | NA
| Analytical results for Co in OREAS 230b (Certified Value 44.4 ppm) | [18 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 45.9 | 43.1833 | 44.7452 | 39 | 40.4 | 43 | 47 | 45.2 | 49.9 | 42.1 | 48.5081 | 44 | 45 | 44 | 42.1 | 46.33 | 44 | 1
| 1 | 2 | 45.6 | 43.4175 | 44.2083 | 41 | 40.8 | 44 | 47.3 | 45 | 49.6 | 43.2 | 49.0171 | 44 | 46.8 | 44 | 42.6 | 43.79 | 44.3 | 26
| 1 | 3 | 45 | 43.9816 | 45.1327 | 44 | 40.1 | 43.8 | 48.5 | 45 | 48.3 | 42.9 | 50.7786 | 43 | 45 | 44 | 42.5 | 46.48 | 44.5 | 16
| 1 | 4 | 45.1 | 43.2099 | 44.1372 | 43 | 40.4 | 43.1 | 47.9 | 45.3 | 50 | 43.6 | 49.8914 | 44 | 45.8 | 45 | 42.9 | 46.83 | 44.5 | 44.397
| 1 | 5 | 45.9 | 43.7187 | 44.1915 | 46 | 41.6 | 43.1 | 47.1 | 45.1 | 49.2 | 43.3 | 46.8799 | 44 | 46.1 | 44 | 41.8 | 46.33 | 44.5 | 24
| 1 | 6 | 45.3 | 43.6313 | 43.8301 | 42 | 41.1 | 42.7 | 50.5 | 44.7 | 48 | 43.9 | 47.6278 | 44 | 44.8 | 44 | 42.4 | 45.9 | 44.6
| Mean |  | 45.4667 | 43.5237 | 44.3742 | 42.5 | 40.7333 | 43.2833 | 48.05 | 45.05 | 49.1667 | 43.1667 | 48.7838 | 43.8333 | 45.5833 | 44.1667 | 42.3833 | 45.9433 | 44.4
| Median |  | 45.45 | 43.5244 | 44.1999 | 42.5 | 40.6 | 43.1 | 47.6 | 45.05 | 49.4 | 43.25 | 48.7626 | 44 | 45.4 | 44 | 42.45 | 46.33 | 44.5
| Std Dev. |  | 0.39328 | 0.3114 | 0.47457 | 2.429 | 0.55015 | 0.50365 | 1.3263 | 0.20736 | 0.84063 | 0.62503 | 1.4351 | 0.40825 | 0.78592 | 0.40825 | 0.38687 | 1.0965 | 0.21909
| Rel.Std.Dev. |  | 0.865% | 0.7155% | 1.069% | 5.715% | 1.351% | 1.164% | 2.76% | 0.4603% | 1.71% | 1.448% | 2.942% | 0.9314% | 1.724% | 0.9243% | 0.9128% | 2.387% | 0.4934%
| PDM3 |  | 2.409% | -1.967% | -0.0515% | -4.273% | -8.252% | -2.509% | 8.228% | 1.471% | 10.74% | -2.771% | 9.881% | -1.27% | 2.672% | -0.5189% | -4.536% | 3.483% | 0.006647%
| Z-Score (Absolute) |  | 0.61 | 0.47 | 0 | 1.05 | 2.03 | 0.61 | 2.05 | 0.38 | 2.68 | 0.67 | 2.46 | 0.3 | 0.68 | 0.12 | 1.11 | 0.88 | 0.01
| Analytical results for Cr in OREAS 230b (Certified Value 86 ppm) | [18 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 102 | 97.5325 | 85.8605 | 59 | 93 | 79 | 82 | 83 | 82 | 84 | 124.9866 | 66 | 100 | 84 | 91.65 | 97 |  | 1
| 1 | 2 | 101 | 97.0265 | 85.2883 | 59 | 95 | 78 | 81 | 82 | 86 | 86 | 138.3541 | 67 | 102 | 87 | 88.11 | 95 |  | 27
| 1 | 3 | 103 | 96.894 | 86.837 | 62 | 94 | 81 | 83 | 83 | 82 | 83 | 116.241 | 70 | 101 | 88 | 90.46 | 97 |  | 16
| 1 | 4 | 104 | 97.958 | 85.5758 | 61 | 95 | 78 | 83 | 82 | 80 | 86 | 142.0388 | 63 | 103 | 92 | 94.63 | 96 |  | 86.2362
| 1 | 5 | 103 | 97.0365 | 83.2663 | 63 | 98 | 76 | 82 | 81 | 83 | 82 | 152.653 | 67 | 101 | 88 | 85.96 | 96 |  | 25
| 1 | 6 | 105 | 95.6445 | 83.8884 | 59 | 97 | 78 | 83 | 82 | 82 | 84 | 120.1219 | 69 | 102 | 94 | 86.64 | 96
| Mean |  | 103 | 97.0153 | 85.1194 | 60.5 | 95.3333 | 78.3333 | 82.3333 | 82.1667 | 82.5 | 84.1667 | 132.3992 | 67 | 101.5 | 88.8333 | 89.575 | 96.1667
| Median |  | 103 | 97.0315 | 85.4321 | 60 | 95 | 78 | 82.5 | 82 | 82 | 84 | 131.6704 | 67 | 101.5 | 88 | 89.285 | 96
| Std Dev. |  | 1.4142 | 0.7811 | 1.318 | 1.7607 | 1.8619 | 1.633 | 0.8165 | 0.75277 | 1.9748 | 1.6021 | 14.18 | 2.4495 | 1.0488 | 3.6009 | 3.3006 | 0.75277
| Rel.Std.Dev. |  | 1.373% | 0.8051% | 1.548% | 2.91% | 1.953% | 2.085% | 0.9917% | 0.9162% | 2.394% | 1.903% | 10.71% | 3.656% | 1.033% | 4.054% | 3.685% | 0.7828%
| PDM3 |  | 19.44% | 12.5% | -1.295% | -29.84% | 10.55% | -9.164% | -4.526% | -4.719% | -4.333% | -2.4% | 53.53% | -22.31% | 17.7% | 3.012% | 3.872% | 11.52%
| Z-Score (Absolute) |  | 1.27 | 0.8 | 0.15 | 2.1 | 0.66 | 0.69 | 0.37 | 0.38 | 0.36 | 0.22 | 3.6 | 1.58 | 1.15 | 0.15 | 0.21 | 0.73
| Analytical results for Cs in OREAS 230b (Certified Value 0.42 ppm) | [18 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.42 | 0.33832 | 0.36636 | 0.34 | 0.42 | 0.43 | 0.41 | 0.41 | 0.46 | 0.42 | 0.96879 | 0.4 | 0.55 | 0.41 | 0.7 |  |  | 1
| 1 | 2 | 0.41 | 0.37333 | 0.39297 | 0.45 | 0.45 | 0.42 | 0.41 | 0.42 | 0.45 | 0.42 | 0.96041 | 0.4 | 0.59 | 0.44 | 0.7 |  |  | 28
| 1 | 3 | 0.42 | 0.34127 | 0.37669 | 0.36 | 0.44 | 0.42 | 0.42 | 0.41 | 0.44 | 0.42 | 0.96962 | 0.4 | 0.57 | 0.44 | 0.8 |  |  | 16
| 1 | 4 | 0.42 | 0.36345 | 0.36256 | 0.4 | 0.49 | 0.41 | 0.41 | 0.4 | 0.46 | 0.42 | 0.96842 | 0.4 | 0.55 | 0.43 | 0.7 |  |  | 0.41853
| 1 | 5 | 0.44 | 0.37561 | 0.3678 | 0.39 | 0.44 | 0.42 | 0.4 | 0.42 | 0.45 | 0.41 | 0.92976 | 0.4 | 0.56 | 0.42 | 0.7 |  |  | 26
| 1 | 6 | 0.41 | 0.3675 | 0.38519 | 0.5 | 0.4 | 0.43 | 0.42 | 0.41 | 0.44 | 0.43 | 0.93639 | 0.4 | 0.57 | 0.43 | 0.7
| Mean |  | 0.42 | 0.35991 | 0.37526 | 0.40667 | 0.44 | 0.42167 | 0.41167 | 0.41167 | 0.45 | 0.42 | 0.95557 | 0.4 | 0.565 | 0.42833 | 0.71667
| Median |  | 0.42 | 0.36548 | 0.37225 | 0.395 | 0.44 | 0.42 | 0.41 | 0.41 | 0.45 | 0.42 | 0.96442 | 0.4 | 0.565 | 0.43 | 0.7
| Std Dev. |  | 0.010954 | 0.016188 | 0.011921 | 0.059217 | 0.030332 | 0.0075277 | 0.0075277 | 0.0075277 | 0.0089443 | 0.0063246 | 0.017857 | 6.0809e-17 | 0.015166 | 0.01169 | 0.040825
| Rel.Std.Dev. |  | 2.608% | 4.498% | 3.177% | 14.56% | 6.894% | 1.785% | 1.829% | 1.829% | 1.988% | 1.506% | 1.869% | 1.52e-14% | 2.684% | 2.729% | 5.696%
| PDM3 |  | 0.3521% | -14% | -10.34% | -2.834% | 5.131% | 0.7503% | -1.639% | -1.639% | 7.52% | 0.3521% | 128.3% | -4.427% | 35% | 2.343% | 71.24%
| Z-Score (Absolute) |  | 0 | 3.04 | 2.26 | 0.67 | 1.01 | 0.08 | 0.42 | 0.42 | 1.52 | 0 | 27.09 | NA | 7.33 | 0.42 | NA
Results from laboratories 18 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 230b (Certified Value 180 ppm) | [18 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 0
|  |  |  | 0.25g | [15 empty] | 0
| 1 | 1 | 178.7 | 189.85 | 180.4669 | 160 | 168 | 181.5 | 183 | 182 | 185 | 182 | 255.6288 | 187 | 171.5 | 174.5 | 156.5 | 182.7 | 170.4 | 1
| 1 | 2 | 176.3 | 191.9 | 179.1865 | 166 | 167 | 180.5 | 181.5 | 180 | 182 | 185 | 235.6186 | 187 | 173.8 | 178.3 | 159.2 | 184.1 | 171.9 | 29
| 1 | 3 | 178.1 | 190.32 | 180.7684 | 179 | 168 | 182 | 183 | 181.5 | 184.5 | 185 | 239.902 | 186 | 172.7 | 175.6 | 160.1 | 185.4 | 171 | 16
| 1 | 4 | 176.7 | 190.61 | 176.8029 | 177 | 168 | 181 | 186.5 | 182 | 181.5 | 184.5 | 243.2139 | 187 | 172.5 | 178.4 | 161 | 187.4 | 172.5 | 179.7329
| 1 | 5 | 180.5 | 191.76 | 177.2792 | 186 | 173 | 180 | 182.5 | 181 | 184 | 186 | 254.6411 | 186 | 175.6 | 176.8 | 158.4 | 181.4 | 172.5 | 27
| 1 | 6 | 175.6 | 190.57 | 175.0473 | 171 | 174 | 179 | 192 | 181 | 177.5 | 186 | 244.3986 | 187 | 170.6 | 179.6 | 160.8 | 185 | 171.3
| Mean |  | 177.65 | 190.835 | 178.2585 | 173.1667 | 169.6667 | 180.6667 | 184.75 | 181.25 | 182.4167 | 184.75 | 245.5672 | 186.6667 | 172.7833 | 177.2 | 159.3333 | 184.3333 | 171.6
| Median |  | 177.4 | 190.59 | 178.2329 | 174 | 168 | 180.75 | 183 | 181.25 | 183 | 185 | 243.8062 | 187 | 172.6 | 177.55 | 159.65 | 184.55 | 171.6
| Std Dev. |  | 1.8086 | 0.81806 | 2.2551 | 9.411 | 3.0111 | 1.0801 | 3.9338 | 0.75829 | 2.7824 | 1.4748 | 8.0193 | 0.5164 | 1.7589 | 1.9194 | 1.699 | 2.1125 | 0.84853
| Rel.Std.Dev. |  | 1.018% | 0.4287% | 1.265% | 5.435% | 1.775% | 0.5979% | 2.129% | 0.4184% | 1.525% | 0.7983% | 3.266% | 0.2766% | 1.018% | 1.083% | 1.066% | 1.146% | 0.4945%
| PDM3 |  | -1.159% | 6.177% | -0.8203% | -3.653% | -5.601% | 0.5195% | 2.791% | 0.8441% | 1.493% | 2.791% | 36.63% | 3.858% | -3.867% | -1.409% | -11.35% | 2.56% | -4.525%
| Z-Score (Absolute) |  | 0.5 | 1.68 | 0.4 | 1.24 | 1.82 | 0 | 0.67 | 0.1 | 0.29 | 0.67 | 10.72 | 0.99 | 1.3 | 0.57 | 3.52 | 0.61 | 1.5
| Analytical results for Dy in OREAS 230b (Certified Value 4.25 ppm) | [18 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 4.3944 | 4.1148 | 4.9 | 4.65 | 3.5801 | 4.04 | [11 empty] | 1
| 1 | 2 | 4.3762 | 4.098 | 4.6 | 4.69 | 3.5327 | 4.15 | [11 empty] | 4
| 1 | 3 | 4.4275 | 4.0704 | 4.6 | 4.57 | 3.6861 | 4.03 | [11 empty] | 16
| 1 | 4 | 4.3475 | 4.0572 | 4.4 | 5.18 | 3.6298 | 4.12 | [11 empty] | 4.2489
| 1 | 5 | 4.4101 | 4.0443 | 4.6 | 4.83 | 3.4651 | 4.1 | [11 empty] | 28
| 1 | 6 | 4.3743 | 4.0329 | 4.6 | 4.61 | 3.4989 | 4.15
| Mean |  | 4.3883 | 4.0696 | 4.6167 | 4.755 | 3.5654 | 4.0983
| Median |  | 4.3853 | 4.0638 | 4.6 | 4.67 | 3.5564 | 4.11
| Std Dev. |  | 0.028473 | 0.031595 | 0.16021 | 0.22661 | 0.083066 | 0.052694
| Rel.Std.Dev. |  | 0.6488% | 0.7764% | 3.47% | 4.766% | 2.33% | 1.286%
| PDM3 |  | 3.282% | -4.22% | 8.656% | 11.91% | -16.09% | -3.544%
| Z-Score (Absolute) |  | 0.36 | 0.43 | 0.92 | 1.26 | 1.67 | 0.36
| Analytical results for Er in OREAS 230b (Certified Value 2.67 ppm) | [18 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 2.7541 | 2.6146 | 2.9 | 2.84 | 2.1925 | 2.55 | [11 empty] | 1
| 1 | 2 | 2.7638 | 2.5827 | 3.1 | 2.62 | 2.1538 | 2.55 | [11 empty] | 5
| 1 | 3 | 2.7807 | 2.4695 | 2.8 | 2.57 | 2.2273 | 2.55 | [11 empty] | 16
| 1 | 4 | 2.7384 | 2.4203 | 2.4 | 2.98 | 2.2377 | 2.54 | [11 empty] | 2.6669
| 1 | 5 | 2.7413 | 2.544 | 2.6 | 2.63 | 2.0611 | 2.69 | [11 empty] | 29
| 1 | 6 | 2.7298 | 2.499 | 2.8 | 2.64 | 2.1167 | 2.61
| Mean |  | 2.7514 | 2.5217 | 2.7667 | 2.7133 | 2.1649 | 2.5817
| Median |  | 2.7477 | 2.5215 | 2.8 | 2.635 | 2.1732 | 2.55
| Std Dev. |  | 0.018755 | 0.072602 | 0.24221 | 0.16046 | 0.068095 | 0.058793
| Rel.Std.Dev. |  | 0.6817% | 2.879% | 8.755% | 5.914% | 3.145% | 2.277%
| PDM3 |  | 3.165% | -5.446% | 3.739% | 1.739% | -18.83% | -3.197%
| Z-Score (Absolute) |  | 0.63 | 0.76 | 0.72 | 0.4 | 2.92 | 0.4
| Analytical results for Eu in OREAS 230b (Certified Value 1 ppm) | [18 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.89616 | 0.96892 | 1.18 | 1.03 | 1.595 | 0.96 | [11 empty] | 1
| 1 | 2 | 0.96425 | 1.0065 | 1.16 | 1.06 | 1.491 | 0.97 | [11 empty] | 6
| 1 | 3 | 0.9102 | 1.0021 | 0.97 | 1.09 | 1.5263 | 0.93 | [11 empty] | 16
| 1 | 4 | 0.96334 | 0.95827 | 1.05 | 1.12 | 1.5299 | 0.95 | [11 empty] | 0.99936
| 1 | 5 | 0.9623 | 0.94677 | 0.95 | 1.06 | 1.5897 | 0.97 | [11 empty] | 30
| 1 | 6 | 0.92734 | 0.94467 | 1.08 | 1.02 | 1.4992 | 0.98
| Mean |  | 0.93727 | 0.97121 | 1.065 | 1.0633 | 1.5385 | 0.96
| Median |  | 0.94482 | 0.9636 | 1.065 | 1.06 | 1.5281 | 0.965
| Std Dev. |  | 0.030185 | 0.027105 | 0.094816 | 0.037238 | 0.044361 | 0.017889
| Rel.Std.Dev. |  | 3.221% | 2.791% | 8.903% | 3.502% | 2.883% | 1.863%
| PDM3 |  | -6.213% | -2.817% | 6.568% | 6.401% | 53.95% | -3.939%
| Z-Score (Absolute) |  | 1.03 | 0.59 | 0.61 | 0.59 | 6.69 | 0.74
| Analytical results for Fe in OREAS 230b (Certified Value 8.96 wt.%) | [18 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 9.22 | 9.2747 | 8.8244 | 9.21 | 8.87 | 9.16 | 9.23 | 8.94 | 8.96 | 8.85 | 11.4973 | 8.89 | 8.06 | 8.7874 | 8.611 | 8.73 | 9.04 | 1
| 1 | 2 | 9.34 | 9.2604 | 8.7998 | 8.98 | 8.68 | 9.06 | 9.32 | 8.81 | 8.96 | 8.98 | 11.1915 | 8.85 | 8.11 | 8.9286 | 8.6993 | 8.82 | 9.05 | 0
| 1 | 3 | 9.31 | 9.2617 | 9.0141 | 9.27 | 8.8 | 9.05 | 9.3 | 8.94 | 8.87 | 9 | 11.2185 | 8.92 | 8.07 | 8.7462 | 8.6673 | 8.69 | 9.02 | 16
| 1 | 4 | 9.23 | 9.267 | 8.7244 | 9.12 | 8.71 | 8.86 | 9.39 | 8.88 | 8.69 | 8.97 | 11.2901 | 8.89 | 8.16 | 8.9261 | 8.6956 | 8.87 | 9.08 | 8.9645
| 1 | 5 | 9.17 | 9.2864 | 8.8039 | 9.1 | 8.44 | 8.9 | 9.15 | 8.85 | 8.77 | 8.97 | 11.8434 | 8.93 | 8.05 | 8.8894 | 8.5373 | 8.87 | 9.02 | 31
| 1 | 6 | 9.31 | 9.2984 | 8.7653 | 9.16 | 8.84 | 8.88 | 9.61 | 8.81 | 8.69 | 8.98 | 11.273 | 8.93 | 8.1 | 8.9777 | 8.6797 | 8.83 | 9.06
| Mean |  | 9.2633 | 9.2747 | 8.822 | 9.14 | 8.7233 | 8.985 | 9.3333 | 8.8717 | 8.8233 | 8.9583 | 11.3856 | 8.9017 | 8.0917 | 8.8759 | 8.6484 | 8.8017 | 9.045
| Median |  | 9.27 | 9.2708 | 8.8018 | 9.14 | 8.755 | 8.975 | 9.31 | 8.865 | 8.82 | 8.975 | 11.2815 | 8.905 | 8.085 | 8.9078 | 8.6735 | 8.825 | 9.045
| Std Dev. |  | 0.066232 | 0.015056 | 0.10045 | 0.0998 | 0.15706 | 0.12194 | 0.15832 | 0.059133 | 0.12485 | 0.054191 | 0.24879 | 0.031252 | 0.040702 | 0.089997 | 0.063083 | 0.074944 | 0.023452
| Rel.Std.Dev. |  | 0.715% | 0.1623% | 1.139% | 1.092% | 1.8% | 1.357% | 1.696% | 0.6665% | 1.415% | 0.6049% | 2.185% | 0.3511% | 0.503% | 1.014% | 0.7294% | 0.8515% | 0.2593%
| PDM3 |  | 3.333% | 3.461% | -1.59% | 1.958% | -2.69% | 0.2286% | 4.114% | -1.036% | -1.575% | -0.06891% | 27.01% | -0.701% | -9.737% | -0.9885% | -3.527% | -1.817% | 0.8979%
| Z-Score (Absolute) |  | 1.7 | 1.76 | 0.37 | 1.12 | 0.84 | 0.39 | 2.03 | 0.14 | 0.37 | 0.27 | 11.69 | 0 | 3.81 | 0.12 | 1.19 | 0.47 | 0.67
| Analytical results for Ga in OREAS 230b (Certified Value 15.8 ppm) | [18 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 15.96 | 16.3095 | 15.3165 | 15.9 | 15.7 | 15.9 | 16.35 | 16.6 | 15.95 | 15 | 17.949 | 15.4 | 17 | 18 | 11 | 14.45 | 16 | 1
| 1 | 2 | 15.96 | 15.8415 | 14.8998 | 15.4 | 15.9 | 15.5 | 16.5 | 16.55 | 15.8 | 15.05 | 18.2357 | 15.1 | 17.4 | 18 | 11 | 14.74 | 16 | 33
| 1 | 3 | 15.75 | 15.432 | 15.43 | 15.4 | 15.8 | 15.7 | 16.7 | 16.9 | 15.3 | 15.1 | 18.3723 | 15 | 17.1 | 18 | 11 | 14.92 | 15 | 16
| 1 | 4 | 15.6 | 15.731 | 15.7607 | 15.5 | 15.3 | 15.6 | 16.7 | 17 | 16.1 | 15.25 | 18.4365 | 15.5 | 17.6 | 18 | 11 | 14.78 | 16 | 15.8493
| 1 | 5 | 16.41 | 16.3095 | 15.4113 | 16.1 | 15.3 | 15.7 | 15.95 | 16.85 | 17.15 | 15.1 | 17.3165 | 15.7 | 17.7 | 18 | 11 | 14.66 | 16 | 32
| 1 | 6 | 15.65 | 16.2835 | 15.4475 | 15.8 | 16.4 | 15.5 | 17.25 | 16.5 | 15.05 | 15.5 | 17.5701 | 15.2 | 16.9 | 18 | 11 | 14.42 | 16
| Mean |  | 15.8883 | 15.9845 | 15.3776 | 15.6833 | 15.7333 | 15.65 | 16.575 | 16.7333 | 15.8917 | 15.1667 | 17.98 | 15.3167 | 17.2833 | 18 | 11 | 14.6617 | 15.8333
| Median |  | 15.855 | 16.0625 | 15.4207 | 15.65 | 15.75 | 15.65 | 16.6 | 16.725 | 15.875 | 15.1 | 18.0924 | 15.3 | 17.25 | 18 | 11 | 14.7 | 16
| Std Dev. |  | 0.29715 | 0.37165 | 0.27838 | 0.29269 | 0.41312 | 0.15166 | 0.43215 | 0.20897 | 0.73445 | 0.18348 | 0.45541 | 0.26394 | 0.33116 | 0 | 0 | 0.19498 | 0.40825
| Rel.Std.Dev. |  | 1.87% | 2.325% | 1.81% | 1.866% | 2.626% | 0.9691% | 2.607% | 1.249% | 4.622% | 1.21% | 2.533% | 1.723% | 1.916% | 0% | 0% | 1.33% | 2.578%
| PDM3 |  | 0.2462% | 0.853% | -2.976% | -1.047% | -0.7317% | -1.258% | 4.579% | 5.578% | 0.2673% | -4.307% | 13.44% | -3.361% | 9.048% | 13.57% | -30.6% | -7.493% | -0.1008%
| Z-Score (Absolute) |  | 0.11 | 0.25 | 0.63 | 0.19 | 0.11 | 0.23 | 1.11 | 1.34 | 0.12 | 0.94 | 3.15 | 0.72 | 2.14 | NA | NA | 1.67 | NA
Results from laboratories 20, 21 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 230b (Certified Value 3.62 ppm) | [18 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 3.741 | 3.4687 | 3.4 | 3.42 | 0.77147 | 3.51 | [11 empty] | 1
| 1 | 2 | 3.7028 | 3.625 | 3.6 | 3.98 | 0.74573 | 3.62 | [11 empty] | 7
| 1 | 3 | 3.7689 | 3.6185 | 3.7 | 3.54 | 0.78504 | 3.44 | [11 empty] | 16
| 1 | 4 | 3.7709 | 3.4577 | 3.5 | 3.7 | 0.76945 | 3.55 | [11 empty] | 3.6156
| 1 | 5 | 3.7583 | 3.6389 | 3.2 | 4.04 | 0.72133 | 3.56 | [11 empty] | 33
| 1 | 6 | 3.7425 | 3.5547 | 3.5 | 3.78 | 0.75129 | 3.58
| Mean |  | 3.7474 | 3.5606 | 3.4833 | 3.7433 | 0.75738 | 3.5433
| Median |  | 3.7504 | 3.5866 | 3.5 | 3.74 | 0.76037 | 3.555
| Std Dev. |  | 0.025246 | 0.080866 | 0.17224 | 0.24213 | 0.022722 | 0.062183
| Rel.Std.Dev. |  | 0.6737% | 2.271% | 4.945% | 6.468% | 3% | 1.755%
| PDM3 |  | 3.645% | -1.521% | -3.658% | 3.533% | -79.05% | -1.999%
| Z-Score (Absolute) |  | 1.01 | 0.04 | 0.36 | 0.99 | 14.5 | 0.04
| Analytical results for Ge in OREAS 230b (Indicative Value 0.12 ppm) | [18 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.20 | 1.24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.4 | < 1 | 0.6 | 0.15 | 0.05 | 0.09 | 0.08 | 0.11 | 0.13 | 1.1 | 0.2 | 1.3 | 3.4 |  |  |  |  | 1
| 1 | 2 | 1.5 | < 1 | 0.6 | 0.13 | 0.05 | 0.09 | 0.07 | 0.12 | 0.1 | 1.1 | 0.2 | 1.2 | 3.3 |  |  |  |  | 1
| 1 | 3 | 1.5 | < 1 | 0.6 | 0.14 | < 0.05 | 0.1 | 0.09 | 0.09 | 0.17 | 1.2 | 0.2 | 1.3 | 3.2 |  |  |  |  | 16
| 1 | 4 | 1.4 | < 1 | 0.8 | 0.15 | 0.05 | 0.1 | 0.08 | 0.12 | 0.23 | 1.2 | 0.2 | 1.3 | 3.9 |  |  |  |  | 0.12167
| 1 | 5 | 1.4 | < 1 | 0.6 | 0.15 | < 0.05 | 0.1 | 0.09 | 0.1 | 0.18 | 1.2 | 0.2 | 1.3 | 3 |  |  |  |  | 7
| 1 | 6 | 1.5 | < 1 | 0.6 | 0.15 | 0.05 | 0.11 | 0.1 | 0.11 | 0.18 | 1.2 | 0.2 | 1.4 | 3.5
| Mean |  | 1.45 |  | 0.63333 | 0.145 | 0.05 | 0.098333 | 0.085 | 0.10833 | 0.165 | 1.1667 | 0.2 | 1.3 | 3.3833
| Median |  | 1.45 |  | 0.6 | 0.15 | 0.05 | 0.1 | 0.085 | 0.11 | 0.175 | 1.2 | 0.2 | 1.3 | 3.35
| Std Dev. |  | 0.054772 |  | 0.08165 | 0.0083666 | 0 | 0.0075277 | 0.010488 | 0.01169 | 0.045056 | 0.05164 | 3.0405e-17 | 0.063246 | 0.30605
| Rel.Std.Dev. |  | 3.777% |  | 12.89% | 5.77% | 0% | 7.655% | 12.34% | 10.79% | 27.31% | 4.426% | 1.52e-14% | 4.865% | 9.046%
| PDM3 |  | 1092% |  | 420.5% | 19.18% | -58.9% | -19.18% | -30.14% | -10.96% | 35.62% | 858.9% | 64.38% | 968.5% | 2681%
| Z-Score (Absolute) |  | 5.32 | 1.28 | 1.85 | 0.23 | 0.67 | 0.43 | 0.49 | 0.39 | 0.15 | 4.12 | 0 | 4.68 | 13.56
| Analytical results for Hf in OREAS 230b (Certified Value 1.78 ppm) | [18 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
|  |  |  | 0.25g | [15 empty] | 3
| 1 | 1 | 1.88 | 0.64107 | 1.7006 | 1.9 | 2.95 | 1.7 | 1.6 | 1.7 | 1.8 | 1.7 | 2.1995 | 1.8 | 2.34 | 1.74 | 2.01 | 2.4 |  | 1
| 1 | 2 | 1.71 | 0.63584 | 1.6853 | 1.7 | 3 | 1.7 | 1.6 | 1.7 | 2 | 1.6 | 2.1451 | 1.8 | 2.33 | 1.83 | 1.94 | 2.6 |  | 19
| 1 | 3 | 1.7 | 0.65477 | 1.74 | 1.8 | 3.34 | 1.7 | 1.6 | 1.7 | 1.7 | 1.6 | 2.0879 | 1.8 | 2.48 | 1.76 | 1.98 | 2.4 |  | 16
| 1 | 4 | 1.73 | 0.65134 | 1.7167 | 1.7 | 2.89 | 1.7 | 1.8 | 1.7 | 1.8 | 1.6 | 2.159 | 1.8 | 2.43 | 1.79 | 2.04 | 2.6 |  | 1.784
| 1 | 5 | 1.8 | 0.64111 | 1.7495 | 1.6 | 2.97 | 1.6 | 1.6 | 1.7 | 1.8 | 1.7 | 2.2339 | 1.8 | 2.47 | 1.73 | 2.09 | 2.1 |  | 34
| 1 | 6 | 1.66 | 0.68801 | 1.648 | 1.9 | 3.16 | 1.6 | 1.7 | 1.7 | 1.7 | 1.6 | 2.1714 | 1.7 | 2.42 | 1.9 | 1.92 | 2.2
| Mean |  | 1.7467 | 0.65202 | 1.7067 | 1.7667 | 3.0517 | 1.6667 | 1.65 | 1.7 | 1.8 | 1.6333 | 2.1661 | 1.7833 | 2.4117 | 1.7917 | 1.9967 | 2.3833
| Median |  | 1.72 | 0.64623 | 1.7086 | 1.75 | 2.985 | 1.7 | 1.6 | 1.7 | 1.8 | 1.6 | 2.1652 | 1.8 | 2.425 | 1.775 | 1.995 | 2.4
| Std Dev. |  | 0.079917 | 0.018997 | 0.037382 | 0.12111 | 0.16774 | 0.05164 | 0.083666 | 0 | 0.10954 | 0.05164 | 0.049679 | 0.040825 | 0.063692 | 0.064317 | 0.063456 | 0.20412
| Rel.Std.Dev. |  | 4.575% | 2.914% | 2.19% | 6.855% | 5.497% | 3.098% | 5.071% | 0% | 6.086% | 3.162% | 2.293% | 2.289% | 2.641% | 3.59% | 3.178% | 8.565%
| PDM3 |  | -2.092% | -63.45% | -4.334% | -0.9707% | 71.06% | -6.576% | -7.51% | -4.708% | 0.8978% | -8.445% | 21.42% | -0.03646% | 35.18% | 0.4307% | 11.92% | 33.6%
| Z-Score (Absolute) |  | 0.16 | 6.49 | 0.39 | 0.05 | 7.38 | 0.63 | 0.72 | 0.43 | 0.14 | 0.82 | 2.26 | 0.05 | 3.68 | 0.1 | 1.28 | 3.52
| Analytical results for Hg in OREAS 230b (Indicative Value 0.068 ppm) | [18 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 1.06 | 1.21 | [15 empty] | ppm
|  |  | 4A*MS | 4A*MS | [15 empty] | 3
3
| 1 | 1 | < 1 | 0.08 | [15 empty] | 1
| 1 | 2 | < 1 | < 0.01 | [15 empty] | 8
| 1 | 3 | < 1 | 0.08 | [15 empty] | 16
| 1 | 4 | < 1 | < 0.01 | [15 empty] | 0.068333
| 1 | 5 | < 1 | 0.16 | [15 empty] | 7
| 1 | 6 | < 1 | 0.08
| Mean |  |  | 0.1
| Median |  |  | 0.08
| Std Dev. |  |  | 0.04
| Rel.Std.Dev. |  |  | 40%
| PDM3 |  |  | 46.34%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 230b (Certified Value 0.91 ppm) | [18 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.0585 | 0.86316 | 0.9 | 0.98 | 0.83022 | 0.86 | [11 empty] | 1
| 1 | 2 | 1.0121 | 0.88031 | 1 | 1 | 0.81742 | 0.88 | [11 empty] | 3
| 1 | 3 | 1.0208 | 0.89279 | 1 | 0.93 | 0.85074 | 0.86 | [11 empty] | 16
| 1 | 4 | 1.0165 | 0.84666 | 0.9 | 1.02 | 0.83851 | 0.87 | [11 empty] | 0.91221
| 1 | 5 | 1.0227 | 0.86644 | 0.9 | 0.97 | 0.77925 | 0.88 | [11 empty] | 35
| 1 | 6 | 1.0054 | 0.84753 | 1 | 0.98 | 0.79745 | 0.89
| Mean |  | 1.0227 | 0.86615 | 0.95 | 0.98 | 0.81893 | 0.87333
| Median |  | 1.0186 | 0.8648 | 0.95 | 0.98 | 0.82382 | 0.875
| Std Dev. |  | 0.018639 | 0.018142 | 0.054772 | 0.030332 | 0.026686 | 0.012111
| Rel.Std.Dev. |  | 1.823% | 2.095% | 5.766% | 3.095% | 3.259% | 1.387%
| PDM3 |  | 12.11% | -5.05% | 4.142% | 7.431% | -10.23% | -4.262%
| Z-Score (Absolute) |  | 1.85 | 0.09 | NA | 1.32 | 0.67 | 0
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 230b (Certified Value 0.08 ppm) | [18 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.08 | < 0.2 | < 0.1 | 0.24 | 0.083 | 0.079 | 0.08 | 0.086 | 0.072 | 0.071363 | 0.08 | 0.08 | 0.07 | < 1 | 0.08 |  |  | 1
| 1 | 2 | 0.07 | < 0.2 | < 0.1 | 0.24 | 0.078 | 0.077 | 0.081 | 0.085 | 0.081 | 0.049948 | 0.08 | 0.09 | 0.08 | < 1 | 0.08 |  |  | 21
| 1 | 3 | 0.08 | < 0.2 | < 0.1 | 0.27 | 0.082 | 0.084 | 0.08 | 0.088 | 0.079 | 0.066879 | 0.08 | 0.08 | 0.08 | < 1 | 0.08 |  |  | 16
| 1 | 4 | 0.08 | < 0.2 | 0.1 | 0.26 | 0.08 | 0.076 | 0.078 | 0.089 | 0.083 | 0.061097 | 0.08 | 0.08 | 0.09 | < 1 | 0.08 |  |  | 0.080185
| 1 | 5 | 0.08 | < 0.2 | < 0.1 | 0.24 | 0.078 | 0.081 | 0.077 | 0.085 | 0.077 | 0.064951 | 0.08 | 0.08 | 0.08 | < 1 | 0.09 |  |  | 36
| 1 | 6 | 0.08 | < 0.2 | < 0.1 | 0.27 | 0.084 | 0.086 | 0.081 | 0.085 | 0.083 | 0.064876 | 0.08 | 0.08 | 0.08 | < 1 | 0.08
| Mean |  | 0.078333 |  | 0.1 | 0.25333 | 0.080833 | 0.0805 | 0.0795 | 0.086333 | 0.079167 | 0.063186 | 0.08 | 0.081667 | 0.08 |  | 0.081667
| Median |  | 0.08 |  | 0.1 | 0.25 | 0.081 | 0.08 | 0.08 | 0.0855 | 0.08 | 0.064913 | 0.08 | 0.08 | 0.08 |  | 0.08
| Std Dev. |  | 0.0040825 |  |  | 0.015055 | 0.0025626 | 0.003937 | 0.0016432 | 0.0017512 | 0.0042151 | 0.0072949 | 0 | 0.0040825 | 0.0063246 |  | 0.0040825
| Rel.Std.Dev. |  | 5.212% |  |  | 5.943% | 3.17% | 4.891% | 2.067% | 2.028% | 5.324% | 11.55% | 0% | 4.999% | 7.906% |  | 4.999%
| PDM3 |  | -2.309% |  | 24.71% | 215.9% | 0.8083% | 0.3926% | -0.8545% | 7.667% | -1.27% | -21.2% | -0.2309% | 1.848% | -0.2309% |  | 1.848%
| Z-Score (Absolute) |  | 1.18 | 9.78 | NA | 87.32 | 0.08 | 0.08 | 0.59 | 2.87 | 0.76 | 8.84 | 0.34 | 0.51 | 0.34 | 212.07 | 0.51
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 230b (Certified Value 0.18 wt.%) | [18 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.1757 | 0.17652 | 0.18441 | 0.19 | 0.19 | 0.17 | 0.18 | 0.17 | 0.17 | 0.17 | 0.21136 | 0.17 | 0.19 | 0.1743 | 0.2 | 0.19 |  | 1
| 1 | 2 | 0.1773 | 0.17579 | 0.18434 | 0.19 | 0.18 | 0.17 | 0.18 | 0.17 | 0.17 | 0.18 | 0.21712 | 0.169 | 0.2 | 0.1761 | 0.21 | 0.19 |  | 0
| 1 | 3 | 0.1767 | 0.17493 | 0.18942 | 0.19 | 0.18 | 0.17 | 0.18 | 0.17 | 0.17 | 0.18 | 0.22504 | 0.173 | 0.19 | 0.1755 | 0.22 | 0.19 |  | 16
| 1 | 4 | 0.1743 | 0.17245 | 0.18117 | 0.19 | 0.18 | 0.17 | 0.18 | 0.17 | 0.17 | 0.18 | 0.22224 | 0.17 | 0.19 | 0.18 | 0.18 | 0.19 |  | 0.18033
| 1 | 5 | 0.1723 | 0.17546 | 0.18386 | 0.19 | 0.18 | 0.17 | 0.18 | 0.17 | 0.17 | 0.18 | 0.20737 | 0.173 | 0.19 | 0.1761 | 0.19 | 0.18 |  | 37
| 1 | 6 | 0.1756 | 0.1727 | 0.18222 | 0.19 | 0.19 | 0.17 | 0.19 | 0.17 | 0.17 | 0.18 | 0.21272 | 0.171 | 0.2 | 0.1789 | 0.21 | 0.19
| Mean |  | 0.17532 | 0.17464 | 0.18423 | 0.19 | 0.18333 | 0.17 | 0.18167 | 0.17 | 0.17 | 0.17833 | 0.21597 | 0.171 | 0.19333 | 0.17682 | 0.20167 | 0.18833
| Median |  | 0.17565 | 0.17519 | 0.1841 | 0.19 | 0.18 | 0.17 | 0.18 | 0.17 | 0.17 | 0.18 | 0.21492 | 0.1705 | 0.19 | 0.1761 | 0.205 | 0.19
| Std Dev. |  | 0.0018005 | 0.0016832 | 0.0028465 | 3.0405e-17 | 0.005164 | 0 | 0.0040825 | 0 | 0 | 0.0040825 | 0.0067604 | 0.0016733 | 0.005164 | 0.0021711 | 0.01472 | 0.0040825
| Rel.Std.Dev. |  | 1.027% | 0.9638% | 1.545% | 1.6e-14% | 2.817% | 0% | 2.247% | 0% | 0% | 2.289% | 3.13% | 0.9785% | 2.671% | 1.228% | 7.299% | 2.168%
| PDM3 |  | -2.782% | -3.158% | 2.163% | 5.36% | 1.663% | -5.73% | 0.7391% | -5.73% | -5.73% | -1.109% | 19.76% | -5.176% | 7.209% | -1.95% | 11.83% | 4.436%
| Z-Score (Absolute) |  | 0.36 | 0.42 | 0.33 | 0.78 | 0.26 | 0.78 | 0.13 | 0.78 | 0.78 | 0.13 | 2.8 | 0.7 | 1.04 | 0.25 | 1.69 | 0.65
| Analytical results for La in OREAS 230b (Certified Value 4.35 ppm) | [18 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 4.21 | 4.3258 | 4.2189 | 4.3 | 4.6 | 4.5 | 4.2 | 4.1 | 4.6 | 4.4 | 5.0419 | 4.3 | 3.84 | 4.78 | 13 |  |  | 1
| 1 | 2 | 4.21 | 4.4298 | 4.2795 | 4.4 | 4.7 | 4.5 | 4.2 | 4 | 4.5 | 4.3 | 5.0169 | 4.5 | 3.97 | 4.9 | 13 |  |  | 22
| 1 | 3 | 4.32 | 4.3173 | 4.278 | 4.4 | 4.5 | 4.6 | 4.2 | 4.1 | 4.4 | 4.4 | 5.2239 | 4.4 | 3.82 | 4.77 | 13 |  |  | 16
| 1 | 4 | 4.28 | 4.4601 | 4.069 | 4.3 | 4.9 | 4.5 | 4.1 | 4.1 | 4.6 | 4.5 | 5.0852 | 4.3 | 3.89 | 4.86 | 13 |  |  | 4.3494
| 1 | 5 | 4.34 | 4.393 | 4.0454 | 4.2 | 4.7 | 4.5 | 4.1 | 4.1 | 4.5 | 4.4 | 4.8397 | 4.5 | 3.95 | 4.83 | 13 |  |  | 38
| 1 | 6 | 4.26 | 4.1865 | 4.0929 | 4.4 | 4.6 | 4.4 | 4.3 | 4 | 4.3 | 4.5 | 4.9302 | 4.3 | 4.01 | 4.72 | 13
| Mean |  | 4.27 | 4.3521 | 4.164 | 4.3333 | 4.6667 | 4.5 | 4.1833 | 4.0667 | 4.4833 | 4.4167 | 5.023 | 4.3833 | 3.9133 | 4.81 | 13
| Median |  | 4.27 | 4.3594 | 4.1559 | 4.35 | 4.65 | 4.5 | 4.2 | 4.1 | 4.5 | 4.4 | 5.0294 | 4.35 | 3.92 | 4.805 | 13
| Std Dev. |  | 0.054406 | 0.098673 | 0.10723 | 0.08165 | 0.13663 | 0.063246 | 0.075277 | 0.05164 | 0.1169 | 0.075277 | 0.1318 | 0.098319 | 0.075542 | 0.065727 | 0
| Rel.Std.Dev. |  | 1.274% | 2.267% | 2.575% | 1.884% | 2.928% | 1.405% | 1.799% | 1.27% | 2.608% | 1.704% | 2.624% | 2.243% | 1.93% | 1.366% | 0%
| PDM3 |  | -1.826% | 0.06081% | -4.264% | -0.3703% | 7.294% | 3.462% | -3.819% | -6.501% | 3.078% | 1.546% | 15.49% | 0.7793% | -10.03% | 10.59% | 198.9%
| Z-Score (Absolute) |  | 0.42 | 0.07 | 0.87 | 0.15 | 1.27 | 0.56 | 0.79 | 1.28 | 0.49 | 0.21 | 2.79 | 0.07 | 1.94 | 1.88 | NA
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 230b (Certified Value 10.1 ppm) | [18 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 9.6 | 10.3892 | 10.6918 | 8 | 12.1 | 10.6 | 10.6 | 10.5 | 10.2 | 10.5 | 9.9656 | 8.9 | 10 | 9.3 | 12.09 | 9.9 |  | 1
| 1 | 2 | 9.5 | 10.0494 | 11.1717 | 8 | 11.9 | 10.6 | 10.7 | 10.2 | 10.2 | 10.5 | 10.3117 | 8.9 | 11 | 9.5 | 11.83 | 9.9 |  | 0
| 1 | 3 | 9.6 | 10.0959 | 10.8763 | 9 | 12.2 | 10.9 | 10.7 | 10.6 | 10.2 | 10.5 | 10.8073 | 8.7 | 10 | 9.3 | 12.19 | 9.8 |  | 16
| 1 | 4 | 9.5 | 10.2408 | 10.8053 | 10 | 12 | 10.4 | 10.8 | 10.4 | 10 | 10.5 | 10.7668 | 9.1 | 11 | 9.3 | 12.64 | 9.8 |  | 10.1106
| 1 | 5 | 9.6 | 10.3482 | 10.7027 | 11 | 12.2 | 10.4 | 10.5 | 10.2 | 10.1 | 10.6 | 9.9686 | 9.3 | 11 | 9.4 | 12.33 | 9.7 |  | 39
| 1 | 6 | 9.5 | 10.1347 | 10.6301 | 9 | 12.4 | 10.5 | 11 | 10.1 | 10 | 10.6 | 9.8093 | 9.1 | 11 | 9.7 | 11.53 | 9.7
| Mean |  | 9.55 | 10.2097 | 10.813 | 9.1667 | 12.1333 | 10.5667 | 10.7167 | 10.3333 | 10.1167 | 10.5333 | 10.2715 | 9 | 10.6667 | 9.4167 | 12.1017 | 9.8
| Median |  | 9.55 | 10.1877 | 10.754 | 9 | 12.15 | 10.55 | 10.7 | 10.3 | 10.15 | 10.5 | 10.1401 | 9 | 11 | 9.35 | 12.14 | 9.8
| Std Dev. |  | 0.054772 | 0.13902 | 0.19651 | 1.169 | 0.17512 | 0.18619 | 0.17224 | 0.19664 | 0.098319 | 0.05164 | 0.43197 | 0.20976 | 0.5164 | 0.16021 | 0.38742 | 0.089443
| Rel.Std.Dev. |  | 0.5735% | 1.362% | 1.817% | 12.75% | 1.443% | 1.762% | 1.607% | 1.903% | 0.9719% | 0.4903% | 4.205% | 2.331% | 4.841% | 1.701% | 3.201% | 0.9127%
| PDM3 |  | -5.545% | 0.9799% | 6.947% | -9.336% | 20.01% | 4.51% | 5.994% | 2.203% | 0.05971% | 4.181% | 1.591% | -10.98% | 5.5% | -6.864% | 19.69% | -3.072%
| Z-Score (Absolute) |  | 1.11 | 0.14 | 0.75 | NA | 2.69 | 0.39 | 0.61 | 0.05 | 0.27 | 0.34 | 0.05 | 1.92 | NA | 1.3 | 2.65 | 0.74
Results from laboratories 7 and 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 230b (Certified Value 0.37 ppm) | [18 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.06 | 1.07 | 1.09 | 1.17 | 1.19 | 1.20 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.36704 | 0.3 | 0.39 | 0.27895 | 0.35 | 0.37 | < 1 | [10 empty] | 1
| 1 | 2 | 0.35892 | 0.4 | 0.38 | 0.2791 | 0.38 | 0.38 | < 1 | [10 empty] | 4
| 1 | 3 | 0.37384 | 0.4 | 0.42 | 0.28012 | 0.37 | 0.38 | < 1 | [10 empty] | 16
| 1 | 4 | 0.35536 | 0.4 | 0.4 | 0.27707 | 0.36 | 0.37 | < 1 | [10 empty] | 0.37419
| 1 | 5 | 0.35731 | 0.3 | 0.41 | 0.26305 | 0.37 | 0.36 | < 1 | [10 empty] | 40
| 1 | 6 | 0.36326 | 0.4 | 0.41 | 0.26622 | 0.37 | 0.38 | < 1
| Mean |  | 0.36262 | 0.36667 | 0.40167 | 0.27409 | 0.36667 | 0.37333
| Median |  | 0.36109 | 0.4 | 0.405 | 0.27801 | 0.37 | 0.375
| Std Dev. |  | 0.0069348 | 0.05164 | 0.01472 | 0.0074544 | 0.010328 | 0.008165
| Rel.Std.Dev. |  | 1.912% | 14.08% | 3.665% | 2.72% | 2.817% | 2.187%
| PDM3 |  | -3.092% | -2.011% | 7.343% | -26.75% | -2.011% | -0.2292%
| Z-Score (Absolute) |  | 0.41 | 0 | 3.54 | 9.36 | 0 | 0.67 | 13.49
| Analytical results for Mg in OREAS 230b (Certified Value 3.82 wt.%) | [18 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.6961 | 3.9582 | 3.7423 | 4.06 | 3.8 | 3.89 | 3.95 | 3.72 | 3.87 | 3.83 | 4.6349 | 3.84 | 3.58 | 3.6833 | 3.6482 | 3.75 | 4.09 | 1
| 1 | 2 | 3.75 | 3.9475 | 3.7144 | 4.03 | 3.65 | 3.83 | 3.97 | 3.68 | 3.86 | 3.88 | 4.5425 | 3.87 | 3.62 | 3.746 | 3.6969 | 3.71 | 4.07 | 0
| 1 | 3 | 3.7464 | 3.9421 | 3.7947 | 3.99 | 3.98 | 3.96 | 3.98 | 3.77 | 3.85 | 3.88 | 4.5973 | 3.89 | 3.59 | 3.7271 | 3.6777 | 3.74 | 4.07 | 16
| 1 | 4 | 3.6991 | 3.9399 | 3.6804 | 3.99 | 3.75 | 3.81 | 4.01 | 3.72 | 3.75 | 3.88 | 4.6308 | 3.87 | 3.65 | 3.7778 | 3.7023 | 3.8 | 4.01 | 3.8181
| 1 | 5 | 3.6752 | 3.9359 | 3.7254 | 4.04 | 3.65 | 3.77 | 3.92 | 3.7 | 3.79 | 3.89 | 4.8256 | 3.89 | 3.61 | 3.7471 | 3.6263 | 3.83 | 4.03 | 41
| 1 | 6 | 3.755 | 3.9244 | 3.7095 | 4.05 | 3.66 | 3.84 | 4.14 | 3.7 | 3.75 | 3.89 | 4.6096 | 3.86 | 3.62 | 3.7972 | 3.6624 | 3.76 | 4
| Mean |  | 3.7203 | 3.9413 | 3.7278 | 4.0267 | 3.7483 | 3.85 | 3.995 | 3.715 | 3.8117 | 3.875 | 4.6401 | 3.87 | 3.6117 | 3.7464 | 3.669 | 3.765 | 4.045
| Median |  | 3.7228 | 3.941 | 3.7199 | 4.035 | 3.705 | 3.835 | 3.975 | 3.71 | 3.82 | 3.88 | 4.6202 | 3.87 | 3.615 | 3.7466 | 3.67 | 3.755 | 4.05
| Std Dev. |  | 0.034166 | 0.011332 | 0.038596 | 0.030111 | 0.12922 | 0.066633 | 0.077136 | 0.030822 | 0.055287 | 0.022583 | 0.096758 | 0.018974 | 0.024833 | 0.039803 | 0.029198 | 0.043243 | 0.036742
| Rel.Std.Dev. |  | 0.9184% | 0.2875% | 1.035% | 0.7478% | 3.447% | 1.731% | 1.931% | 0.8297% | 1.45% | 0.5828% | 2.085% | 0.4903% | 0.6876% | 1.062% | 0.7958% | 1.149% | 0.9083%
| PDM3 |  | -2.561% | 3.228% | -2.365% | 5.463% | -1.827% | 0.8363% | 4.634% | -2.7% | -0.1677% | 1.491% | 21.53% | 1.36% | -5.406% | -1.877% | -3.905% | -1.39% | 5.944%
| Z-Score (Absolute) |  | 0.67 | 0.96 | 0.62 | 1.59 | 0.47 | 0.28 | 1.35 | 0.71 | 0 | 0.47 | 6.11 | 0.43 | 1.48 | 0.48 | 1.05 | 0.34 | 1.72
| Analytical results for Mn in OREAS 230b (Certified Value 0.151 wt.%) | [18 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.1534 | 0.15466 | 0.14942 | 0.158 | 0.1456 | 0.146 | 0.153 | 0.1455 | 0.1505 | 0.1475 | 0.1869 | 0.153 | 0.1627 | 0.154 | 0.1335 | 0.1383 | 0.1624 | 1
| 1 | 2 | 0.1552 | 0.15509 | 0.14894 | 0.156 | 0.1449 | 0.1445 | 0.151 | 0.144 | 0.1505 | 0.151 | 0.1907 | 0.153 | 0.1651 | 0.1558 | 0.1344 | 0.1395 | 0.1639 | 24
| 1 | 3 | 0.1555 | 0.15592 | 0.15179 | 0.156 | 0.1429 | 0.1475 | 0.1515 | 0.15 | 0.148 | 0.151 | 0.19687 | 0.151 | 0.1604 | 0.1524 | 0.1343 | 0.1378 | 0.1645 | 16
| 1 | 4 | 0.1529 | 0.1554 | 0.14755 | 0.158 | 0.1429 | 0.1445 | 0.154 | 0.145 | 0.146 | 0.1485 | 0.19434 | 0.154 | 0.1623 | 0.1557 | 0.1342 | 0.1404 | 0.1641 | 0.1506
| 1 | 5 | 0.1525 | 0.15582 | 0.14897 | 0.157 | 0.1472 | 0.143 | 0.1495 | 0.145 | 0.1455 | 0.152 | 0.18238 | 0.155 | 0.163 | 0.154 | 0.132 | 0.1407 | 0.1643 | 42
| 1 | 6 | 0.1557 | 0.15421 | 0.14868 | 0.159 | 0.1474 | 0.144 | 0.1585 | 0.146 | 0.1455 | 0.151 | 0.18521 | 0.154 | 0.16 | 0.1561 | 0.1335 | 0.1393 | 0.1634
| Mean |  | 0.1542 | 0.15518 | 0.14922 | 0.15733 | 0.14515 | 0.14492 | 0.15292 | 0.14592 | 0.14767 | 0.15017 | 0.1894 | 0.15333 | 0.16225 | 0.15467 | 0.13365 | 0.13933 | 0.16377
| Median |  | 0.1543 | 0.15524 | 0.14895 | 0.1575 | 0.14525 | 0.1445 | 0.15225 | 0.14525 | 0.147 | 0.151 | 0.1888 | 0.1535 | 0.1625 | 0.15485 | 0.13385 | 0.1394 | 0.164
| Std Dev. |  | 0.0014255 | 0.00066945 | 0.0014033 | 0.0012111 | 0.0019827 | 0.0015943 | 0.003153 | 0.0021075 | 0.0023805 | 0.0017512 | 0.0055668 | 0.0013663 | 0.0018641 | 0.0014445 | 0.00090056 | 0.0011361 | 0.00076855
| Rel.Std.Dev. |  | 0.9244% | 0.4314% | 0.9404% | 0.7697% | 1.366% | 1.1% | 2.062% | 1.444% | 1.612% | 1.166% | 2.939% | 0.891% | 1.149% | 0.934% | 0.6738% | 0.8154% | 0.4693%
| PDM3 |  | 2.387% | 3.04% | -0.9165% | 4.468% | -3.622% | -3.777% | 1.535% | -3.113% | -1.951% | -0.2908% | 25.76% | 1.812% | 7.732% | 2.697% | -11.26% | -7.484% | 8.739%
| Z-Score (Absolute) |  | 0.16 | 0.29 | 0.47 | 0.57 | 1 | 1.03 | 0 | 0.9 | 0.67 | 0.35 | 4.69 | 0.05 | 1.2 | 0.22 | 2.47 | 1.74 | 1.39
| Analytical results for Mo in OREAS 230b (Certified Value 0.57 ppm) | [18 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
2
| 1 | 1 | 0.5 | < 10 | < 50 | < 1 | 1.51 | 0.47 | 0.58 | 0.57 | 0.45 | 0.5 | 3.1976 | 0.5 | 0.78 | 0.5 | 2.9 | < 0.3 | 0.7 | 1
| 1 | 2 | 0.7 | < 10 | < 50 | < 1 | 1.27 | 0.47 | 0.6 | 0.55 | 0.45 | 0.5 | 2.8821 | 0.5 | 0.77 | 0.6 | 2.7 | < 0.3 | 0.5 | 25
| 1 | 3 | 0.6 | < 10 | < 50 | < 1 | 1.48 | 0.54 | 0.55 | 0.55 | 0.53 | 0.52 | 4.0372 | 0.5 | 0.81 | 0.6 | 2.8 | < 0.3 | 0.6 | 16
| 1 | 4 | 0.7 | < 10 | < 50 | < 1 | 1.35 | 0.44 | 0.57 | 0.61 | 0.49 | 0.51 | 4.3562 | 0.5 | 0.82 | 0.5 | 2.6 | < 0.3 | 0.6 | 0.56889
| 1 | 5 | 0.6 | < 10 | < 50 | < 1 | 1.24 | 0.53 | 0.57 | 0.55 | 0.49 | 0.51 | 3.2727 | 0.5 | 0.78 | 0.6 | 2.5 | < 0.3 | 0.7 | 43
| 1 | 6 | 0.5 | < 10 | < 50 | < 1 | 1.32 | 0.45 | 0.62 | 0.62 | 0.44 | 0.48 | 2.9104 | 0.5 | 0.81 | 0.5 | 2.8 | < 0.3 | 0.6
| Mean |  | 0.6 |  |  |  | 1.3617 | 0.48333 | 0.58167 | 0.575 | 0.475 | 0.50333 | 3.4427 | 0.5 | 0.795 | 0.55 | 2.7167 |  | 0.61667
| Median |  | 0.6 |  |  |  | 1.335 | 0.47 | 0.575 | 0.56 | 0.47 | 0.505 | 3.2352 | 0.5 | 0.795 | 0.55 | 2.75 |  | 0.6
| Std Dev. |  | 0.089443 |  |  |  | 0.11053 | 0.041793 | 0.024833 | 0.032094 | 0.034496 | 0.013663 | 0.61228 | 0 | 0.020736 | 0.054772 | 0.1472 |  | 0.075277
| Rel.Std.Dev. |  | 14.91% |  |  |  | 8.117% | 8.647% | 4.269% | 5.581% | 7.262% | 2.714% | 17.78% | 0% | 2.608% | 9.959% | 5.418% |  | 12.21%
| PDM3 |  | 5.469% |  |  |  | 139.4% | -15.04% | 2.246% | 1.074% | -16.5% | -11.52% | 505.2% | -12.11% | 39.75% | -3.32% | 377.5% |  | 8.398%
| Z-Score (Absolute) |  | NA | 18.63 | 102.92 | 0.33 | 3.3 | 0.4 | 0.01 | 0.01 | 0.44 | 0.32 | 12.07 | NA | 0.91 | NA | NA | 1.81 | NA
Results from laboratories 1, 18, 20, 21 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 230b (Certified Value 2 wt.%) | [18 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.9901 | 2.06 | 1.9835 | 2.02 | 1.94 | 2 | 2 | 1.94 | 1.99 | 1.98 | 2.5949 | 2 | 2.27 | 1.9476 | 2.1029 | 1.82 | 2.04 | 1
| 1 | 2 | 2.0149 | 2.0372 | 1.9668 | 2.04 | 1.91 | 1.97 | 2.03 | 1.9 | 1.98 | 2 | 2.5379 | 2 | 2.28 | 1.9731 | 2.1267 | 1.75 | 2.01 | 0
| 1 | 3 | 2.0059 | 2.052 | 2.0146 | 2.04 | 1.92 | 1.98 | 2.02 | 1.95 | 1.98 | 2 | 2.5691 | 2 | 2.29 | 1.9574 | 2.1169 | 1.72 | 2.03 | 16
| 1 | 4 | 1.9806 | 2.0319 | 1.9733 | 2.06 | 1.9 | 1.94 | 2.04 | 1.91 | 1.94 | 2.01 | 2.5842 | 1.99 | 2.31 | 1.9777 | 2.1226 | 1.84 | 2.01 | 1.9961
| 1 | 5 | 1.982 | 2.0551 | 1.9658 | 2.03 | 1.84 | 1.96 | 2 | 1.92 | 1.95 | 2.01 | 2.6918 | 2 | 2.31 | 1.9761 | 2.0904 | 1.85 | 2.01 | 44
| 1 | 6 | 2.0114 | 2.0516 | 1.9584 | 2.03 | 1.93 | 1.93 | 2.11 | 1.91 | 1.93 | 2 | 2.5667 | 2 | 2.28 | 2.0238 | 2.1075 | 1.72 | 2
| Mean |  | 1.9975 | 2.048 | 1.9771 | 2.0367 | 1.9067 | 1.9633 | 2.0333 | 1.9217 | 1.9617 | 2 | 2.5908 | 1.9983 | 2.29 | 1.976 | 2.1112 | 1.7833 | 2.0167
| Median |  | 1.998 | 2.0518 | 1.97 | 2.035 | 1.915 | 1.965 | 2.025 | 1.915 | 1.965 | 2 | 2.5767 | 2 | 2.285 | 1.9746 | 2.1122 | 1.785 | 2.01
| Std Dev. |  | 0.015147 | 0.010967 | 0.020235 | 0.013663 | 0.03559 | 0.02582 | 0.040825 | 0.019408 | 0.024833 | 0.010954 | 0.053125 | 0.0040825 | 0.016733 | 0.026263 | 0.013549 | 0.060222 | 0.015055
| Rel.Std.Dev. |  | 0.7583% | 0.5355% | 1.023% | 0.6708% | 1.867% | 1.315% | 2.008% | 1.01% | 1.266% | 0.5477% | 2.051% | 0.2043% | 0.7307% | 1.329% | 0.6418% | 3.377% | 0.7466%
| PDM3 |  | 0.0672% | 2.596% | -0.9547% | 2.03% | -4.482% | -1.644% | 1.863% | -3.731% | -1.727% | 0.1933% | 29.79% | 0.1098% | 14.72% | -1.012% | 5.762% | -10.66% | 1.028%
| Z-Score (Absolute) |  | 0.02 | 0.91 | 0.39 | 0.7 | 1.69 | 0.64 | 0.64 | 1.41 | 0.67 | 0.03 | 10.9 | 0 | 5.36 | 0.41 | 2.08 | 3.95 | 0.34
| Analytical results for Nb in OREAS 230b (Certified Value 3.75 ppm) | [18 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.75 | 3.7253 | 3.6 | 7.1 | 3.6 | 3.9 | 3.8 | 3.7 | 3.5 | 3.7842 | 3.9 | 3.8 | 3.65 | 3.66 | 3.9 |  |  | 1
| 1 | 2 | 3.66 | 3.6525 | 3.6 | 7.2 | 3.5 | 3.9 | 3.7 | 3.7 | 3.6 | 3.6149 | 3.6 | 4.1 | 3.65 | 3.66 | 3.9 |  |  | 5
| 1 | 3 | 3.69 | 3.65 | 3.7 | 6.8 | 3.5 | 4 | 3.7 | 3.5 | 3.6 | 4.0155 | 3.8 | 3.9 | 3.69 | 3.7 | 3.9 |  |  | 16
| 1 | 4 | 3.69 | 4.0038 | 3.5 | 7 | 3.4 | 4 | 3.9 | 3.6 | 3.7 | 3.9595 | 3.8 | 4.2 | 3.75 | 3.7 | 3.8 |  |  | 3.7505
| 1 | 5 | 3.75 | 3.9416 | 3.5 | 6.6 | 3.5 | 3.8 | 3.8 | 3.7 | 3.6 | 5.2493 | 3.8 | 4.1 | 3.6 | 3.74 | 4 |  |  | 45
| 1 | 6 | 3.68 | 3.6934 | 3.7 | 6.9 | 3.5 | 4.1 | 3.8 | 3.6 | 3.6 | 4.1736 | 3.7 | 4 | 3.73 | 3.77 | 3.8
| Mean |  | 3.7033 | 3.7778 | 3.6 | 6.9333 | 3.5 | 3.95 | 3.7833 | 3.6333 | 3.6 | 4.1328 | 3.7667 | 4.0167 | 3.6783 | 3.705 | 3.8833
| Median |  | 3.69 | 3.7093 | 3.6 | 6.95 | 3.5 | 3.95 | 3.8 | 3.65 | 3.6 | 3.9875 | 3.8 | 4.05 | 3.67 | 3.7 | 3.9
| Std Dev. |  | 0.037771 | 0.15478 | 0.089443 | 0.21602 | 0.063246 | 0.10488 | 0.075277 | 0.08165 | 0.063246 | 0.57996 | 0.10328 | 0.1472 | 0.056006 | 0.043704 | 0.075277
| Rel.Std.Dev. |  | 1.02% | 4.097% | 2.485% | 3.116% | 1.807% | 2.655% | 1.99% | 2.247% | 1.757% | 14.03% | 2.742% | 3.665% | 1.523% | 1.18% | 1.938%
| PDM3 |  | -1.258% | 0.7263% | -4.013% | 84.86% | -6.68% | 5.319% | 0.8749% | -3.125% | -4.013% | 10.19% | 0.4305% | 7.096% | -1.925% | -1.214% | 3.541%
| Z-Score (Absolute) |  | 0.32 | 0.06 | 0.84 | 16.01 | 1.35 | 0.93 | 0.08 | 0.67 | 0.84 | 1.85 | 0 | 1.26 | 0.45 | 0.31 | 0.59
| Analytical results for Nd in OREAS 230b (Certified Value 8.67 ppm) | [18 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | 1.21 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 9.3476 | 8.1681 | 8.3 | 9 | 9.2065 | 7.96 | 11 | [10 empty] | 1
| 1 | 2 | 9.3293 | 8.4114 | 8.1 | 8.8 | 9.2697 | 8.23 | 5 | [10 empty] | 6
| 1 | 3 | 9.3489 | 8.4502 | 8.5 | 8.7 | 9.4846 | 8.17 | 4 | [10 empty] | 16
| 1 | 4 | 9.3828 | 7.8759 | 7.9 | 9.7 | 9.3129 | 8.21 | 6 | [10 empty] | 8.6714
| 1 | 5 | 9.3638 | 7.9032 | 7.9 | 8.6 | 8.9437 | 8.26 | 9 | [10 empty] | 46
| 1 | 6 | 9.3327 | 7.8796 | 8.5 | 8.7 | 9.1811 | 8.45 | 8
| Mean |  | 9.3508 | 8.1147 | 8.2 | 8.9167 | 9.2331 | 8.2133 | 7.1667
| Median |  | 9.3482 | 8.0356 | 8.2 | 8.75 | 9.2381 | 8.22 | 7
| Std Dev. |  | 0.019982 | 0.26852 | 0.27568 | 0.40702 | 0.17782 | 0.15782 | 2.6394
| Rel.Std.Dev. |  | 0.2137% | 3.309% | 3.362% | 4.565% | 1.926% | 1.921% | 36.83%
| PDM3 |  | 7.835% | -6.42% | -5.437% | 2.828% | 6.477% | -5.283% | -17.35%
| Z-Score (Absolute) |  | 1.3 | 0.74 | 0.6 | 0.58 | 1.11 | 0.58 | NA
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 230b (Certified Value 68 ppm) | [18 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 0
1
| 1 | 1 | 69.3 | 64.83 | 72.0435 | 61 | 59 | 69.5 | 74.6 | 71.3 | 72.8 | 64.2 | 88.9691 | 68 | 72.8 | 68.3 | 63.7 | 57.98 | 69.4 | 1
| 1 | 2 | 68.8 | 64.13 | 70.7904 | 65 | 58 | 68.1 | 73.9 | 71.1 | 72 | 65.9 | 82.331 | 68 | 73 | 70.4 | 64.4 | 58.03 | 66.5 | 28
| 1 | 3 | 68.5 | 64.7 | 71.8795 | 68 | 60 | 69.4 | 75.5 | 72.1 | 69.8 | 65.6 | 80.6626 | 69 | 70.9 | 68 | 64 | 58.05 | 67.7 | 16
| 1 | 4 | 68.3 | 63.18 | 71.6387 | 70 | 57 | 68.9 | 74.3 | 72.1 | 69.8 | 66.2 | 90.8273 | 69 | 72.1 | 70.3 | 64.6 | 58.13 | 66.8 | 68.2634
| 1 | 5 | 70.4 | 64.48 | 71.4308 | 72 | 62 | 69.1 | 71.5 | 72.1 | 72.8 | 66.2 | 84.6934 | 69 | 71.8 | 68.9 | 63.2 | 57.23 | 66.7 | 47
| 1 | 6 | 68.5 | 63.29 | 71.8099 | 64 | 59 | 68.5 | 77.9 | 71 | 70.7 | 66.5 | 81.0049 | 69 | 70.5 | 70.7 | 63.4 | 57.38 | 67.2
| Mean |  | 68.9667 | 64.1017 | 71.5988 | 66.6667 | 59.1667 | 68.9167 | 74.6167 | 71.6167 | 71.3167 | 65.7667 | 84.748 | 68.6667 | 71.85 | 69.4333 | 63.8833 | 57.8 | 67.3833
| Median |  | 68.65 | 64.305 | 71.7243 | 66.5 | 59 | 69 | 74.45 | 71.7 | 71.35 | 66.05 | 83.5122 | 69 | 71.95 | 69.6 | 63.85 | 58.005 | 67
| Std Dev. |  | 0.78401 | 0.71284 | 0.44819 | 4.0825 | 1.7224 | 0.53821 | 2.0923 | 0.53821 | 1.4034 | 0.82624 | 4.2737 | 0.5164 | 1.0015 | 1.1759 | 0.55287 | 0.38936 | 1.0759
| Rel.Std.Dev. |  | 1.137% | 1.112% | 0.626% | 6.124% | 2.911% | 0.781% | 2.804% | 0.7515% | 1.968% | 1.256% | 5.043% | 0.752% | 1.394% | 1.694% | 0.8654% | 0.6736% | 1.597%
| PDM3 |  | 1.03% | -6.097% | 4.886% | -2.339% | -13.33% | 0.957% | 9.307% | 4.912% | 4.473% | -3.657% | 24.15% | 0.5908% | 5.254% | 1.714% | -6.416% | -15.33% | -1.289%
| Z-Score (Absolute) |  | 0.01 | 1.2 | 0.67 | 0.56 | 2.44 | 0 | 1.42 | 0.67 | 0.6 | 0.79 | 3.95 | 0.06 | 0.73 | 0.13 | 1.26 | 2.78 | 0.38
| Analytical results for P in OREAS 230b (Certified Value 0.052 wt.%) | [18 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | wt.%
|  |  | 4A*OES | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 3
|  |  |  | 0.25g | [15 empty] | 3
| 1 | 1 | 0.0513 | 0.059129 | 0.049337 | 0.047 | 0.0511 | 0.052 | 0.052 | 0.051 | 0.052 | 0.052 | 0.064185 | 0.052 | 0.05 | 0.0515 | 0.0484 | 0.051 | 0.053 | 1
| 1 | 2 | 0.052 | 0.059367 | 0.048394 | 0.047 | 0.0521 | 0.051 | 0.053 | 0.051 | 0.052 | 0.052 | 0.073864 | 0.052 | 0.053 | 0.0528 | 0.0495 | 0.051 | 0.053 | 0
| 1 | 3 | 0.0519 | 0.059113 | 0.049508 | 0.049 | 0.051 | 0.052 | 0.053 | 0.052 | 0.051 | 0.053 | 0.076555 | 0.053 | 0.0514 | 0.0517 | 0.0487 | 0.053 | 0.053 | 16
| 1 | 4 | 0.0511 | 0.059449 | 0.047648 | 0.049 | 0.0503 | 0.05 | 0.053 | 0.052 | 0.05 | 0.053 | 0.076445 | 0.053 | 0.0516 | 0.0526 | 0.0485 | 0.052 | 0.053 | 0.051655
| 1 | 5 | 0.0509 | 0.059633 | 0.048851 | 0.052 | 0.0506 | 0.051 | 0.052 | 0.052 | 0.051 | 0.053 | 0.071838 | 0.052 | 0.0509 | 0.0522 | 0.0482 | 0.052 | 0.053 | 48
| 1 | 6 | 0.0515 | 0.068933 | 0.048457 | 0.047 | 0.0541 | 0.051 | 0.055 | 0.05 | 0.05 | 0.053 | 0.074024 | 0.052 | 0.0513 | 0.0527 | 0.0481 | 0.051 | 0.052
| Mean |  | 0.05145 | 0.060937 | 0.048699 | 0.0485 | 0.051533 | 0.051167 | 0.053 | 0.051333 | 0.051 | 0.052667 | 0.072818 | 0.052333 | 0.051367 | 0.05225 | 0.048567 | 0.051667 | 0.052833
| Median |  | 0.0514 | 0.059408 | 0.048654 | 0.048 | 0.05105 | 0.051 | 0.053 | 0.0515 | 0.051 | 0.053 | 0.073944 | 0.052 | 0.05135 | 0.0524 | 0.04845 | 0.0515 | 0.053
| Std Dev. |  | 0.00043704 | 0.0039221 | 0.00068422 | 0.0019748 | 0.0013981 | 0.00075277 | 0.0010954 | 0.0008165 | 0.00089443 | 0.0005164 | 0.0045861 | 0.0005164 | 0.00098116 | 0.00054681 | 0.00050465 | 0.0008165 | 0.00040825
| Rel.Std.Dev. |  | 0.8494% | 6.436% | 1.405% | 4.072% | 2.713% | 1.471% | 2.067% | 1.591% | 1.754% | 0.9805% | 6.298% | 0.9867% | 1.91% | 1.047% | 1.039% | 1.58% | 0.7727%
| PDM3 |  | -0.396% | 17.97% | -5.721% | -6.107% | -0.2347% | -0.9445% | 2.605% | -0.6219% | -1.267% | 1.959% | 40.97% | 1.314% | -0.5574% | 1.153% | -5.978% | 0.02343% | 2.282%
| Z-Score (Absolute) |  | 0.07 | 7.93 | 2.39 | 2.56 | 0 | 0.31 | 1.24 | 0.17 | 0.45 | 0.96 | 17.94 | 0.67 | 0.14 | 0.6 | 2.5 | 0.11 | 1.1
| Analytical results for Pb in OREAS 230b (Certified Value 9.59 ppm) | [18 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  |  | 0.25g | [15 empty] | 2
| 1 | 1 | 9.7 | 8.154 | 12.1609 | 9 | 10.5 | 8.5 | 8.4 | 9.1 | 9.4 | 8.2 | 11.6479 | 9 | 9.8 | 9.9 | 16 | 10.5 | 8.5 | 1
| 1 | 2 | 9.4 | 8.168 | 11.3714 | 9 | 10.6 | 8.4 | 8.7 | 9 | 9.3 | 8.6 | 10.8552 | 9 | 9.7 | 10 | 16 | 11.3 | 8.1 | 29
| 1 | 3 | 9.4 | 8.627 | 12.0892 | 9 | 10.3 | 9 | 8.4 | 9.1 | 9 | 8.6 | 10.8859 | 9 | 9.7 | 10.1 | 16 | 11.6 | 8 | 16
| 1 | 4 | 9.2 | 7.494 | 11.0447 | 8 | 10.7 | 8.4 | 8.4 | 8.9 | 9.1 | 8.4 | 12.9673 | 9 | 10.4 | 10.3 | 15 | 11.3 | 8.4 | 9.5883
| 1 | 5 | 9 | 7.648 | 13.2086 | 10 | 10.8 | 8.5 | 8.2 | 9.1 | 9 | 8.4 | 9.9675 | 9 | 10.5 | 10.1 | 15 | 10.9 | 8.3 | 49
| 1 | 6 | 8.9 | 8.121 | 11.1522 | 8 | 11.9 | 8.5 | 9.3 | 9 | 8.8 | 8.4 | 10.2524 | 9 | 9.9 | 10.2 | 17 | 11.7 | 7.9
| Mean |  | 9.2667 | 8.0353 | 11.8378 | 8.8333 | 10.8 | 8.55 | 8.5667 | 9.0333 | 9.1 | 8.4333 | 11.096 | 9 | 10 | 10.1 | 15.8333 | 11.2167 | 8.2
| Median |  | 9.3 | 8.1375 | 11.7303 | 9 | 10.65 | 8.5 | 8.4 | 9.05 | 9.05 | 8.4 | 10.8705 | 9 | 9.85 | 10.1 | 16 | 11.3 | 8.2
| Std Dev. |  | 0.29439 | 0.40796 | 0.82015 | 0.75277 | 0.56569 | 0.22583 | 0.39328 | 0.08165 | 0.21909 | 0.15055 | 1.0857 | 0 | 0.35777 | 0.14142 | 0.75277 | 0.44907 | 0.23664
| Rel.Std.Dev. |  | 3.177% | 5.077% | 6.928% | 8.522% | 5.238% | 2.641% | 4.591% | 0.9039% | 2.408% | 1.785% | 9.785% | 0% | 3.578% | 1.4% | 4.754% | 4.004% | 2.886%
| PDM3 |  | -3.354% | -16.2% | 23.46% | -7.874% | 12.64% | -10.83% | -10.65% | -5.788% | -5.092% | -12.05% | 15.73% | -6.135% | 4.294% | 5.337% | 65.13% | 16.98% | -14.48%
| Z-Score (Absolute) |  | 0 | 1 | 2.08 | NA | 1.24 | 0.58 | 0.57 | 0.19 | 0.13 | 0.67 | 1.48 | NA | 0.59 | 0.67 | 5.31 | 1.58 | 0.86
Results from laboratories 7 and 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 230b (Certified Value 1.65 ppm) | [18 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.17 | 1.20 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.6097 | 1.6405 | 1.6 | 1.68 | 1.862 | 1.63 | [11 empty] | 1
| 1 | 2 | 1.5904 | 1.6844 | 1.7 | 1.65 | 1.8274 | 1.67 | [11 empty] | 7
| 1 | 3 | 1.5967 | 1.6219 | 1.7 | 1.74 | 1.9051 | 1.62 | [11 empty] | 16
| 1 | 4 | 1.6052 | 1.6213 | 1.7 | 1.79 | 1.8865 | 1.61 | [11 empty] | 1.6503
| 1 | 5 | 1.5689 | 1.6403 | 1.7 | 1.63 | 1.7647 | 1.66 | [11 empty] | 50
| 1 | 6 | 1.5802 | 1.6158 | 1.7 | 1.76 | 1.7958 | 1.65
| Mean |  | 1.5919 | 1.6374 | 1.6833 | 1.7083 | 1.8402 | 1.64
| Median |  | 1.5936 | 1.6311 | 1.7 | 1.71 | 1.8447 | 1.64
| Std Dev. |  | 0.01539 | 0.02527 | 0.040825 | 0.064317 | 0.054157 | 0.023664
| Rel.Std.Dev. |  | 0.9668% | 1.543% | 2.425% | 3.765% | 2.943% | 1.443%
| PDM3 |  | -3.541% | -0.782% | 2.001% | 3.516% | 11.51% | -0.6243%
| Z-Score (Absolute) |  | 1.33 | 0.46 | 0.41 | 0.89 | 3.4 | 0.41
| Analytical results for Rb in OREAS 230b (Certified Value 4.64 ppm) | [18 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 4.83 | 4.9387 | 4.7371 | 4.6 | 4.5 | 4.5 | 4.8 | 3.9 | 4.9 | 4.4 | 5.9346 | 4.1 | 5.2 | 4.68 | 4.75 | 7.3 |  | 1
| 1 | 2 | 4.51 | 4.1844 | 4.6432 | 4.1 | 4.3 | 4.3 | 4.9 | 3.8 | 4.9 | 4.4 | 5.7448 | 4.2 | 5.6 | 4.76 | 4.77 | 6.7 |  | 31
| 1 | 3 | 4.71 | 4.7569 | 4.5946 | 4.3 | 4.4 | 4.3 | 4.9 | 3.8 | 4.7 | 4.4 | 5.674 | 4.1 | 5.3 | 4.5 | 4.85 | 7.2 |  | 16
| 1 | 4 | 4.72 | 4.1924 | 4.9165 | 5.2 | 4.6 | 4.3 | 4.9 | 3.9 | 4.9 | 4.5 | 5.8116 | 4.1 | 5.1 | 4.65 | 4.75 | 7.8 |  | 4.6408
| 1 | 5 | 4.69 | 4.2205 | 4.5713 | 4.7 | 4.5 | 4.4 | 4.7 | 3.9 | 4.8 | 4.4 | 6.0407 | 4.3 | 5.4 | 4.74 | 4.76 | 7 |  | 51
| 1 | 6 | 4.55 | 4.9615 | 4.7658 | 4.5 | 4.5 | 4.3 | 5.2 | 3.9 | 4.7 | 4.5 | 5.8066 | 4.1 | 5.7 | 4.6 | 4.66 | 7
| Mean |  | 4.6683 | 4.5424 | 4.7047 | 4.5667 | 4.4667 | 4.35 | 4.9 | 3.8667 | 4.8167 | 4.4333 | 5.8354 | 4.15 | 5.3833 | 4.655 | 4.7567 | 7.1667
| Median |  | 4.7 | 4.4887 | 4.6901 | 4.55 | 4.5 | 4.3 | 4.9 | 3.9 | 4.85 | 4.4 | 5.8091 | 4.1 | 5.35 | 4.665 | 4.755 | 7.1
| Std Dev. |  | 0.11839 | 0.3829 | 0.12912 | 0.37771 | 0.10328 | 0.083666 | 0.16733 | 0.05164 | 0.098319 | 0.05164 | 0.13239 | 0.083666 | 0.23166 | 0.095864 | 0.060553 | 0.37238
| Rel.Std.Dev. |  | 2.536% | 8.429% | 2.744% | 8.271% | 2.312% | 1.923% | 3.415% | 1.336% | 2.041% | 1.165% | 2.269% | 2.016% | 4.303% | 2.059% | 1.273% | 5.196%
| PDM3 |  | 0.5932% | -2.121% | 1.378% | -1.598% | -3.752% | -6.266% | 5.585% | -16.68% | 3.789% | -4.471% | 25.74% | -10.58% | 16% | 0.3059% | 2.497% | 54.43%
| Z-Score (Absolute) |  | 0.02 | 0.38 | 0.14 | 0.3 | 0.62 | 0.99 | 0.76 | 2.53 | 0.49 | 0.73 | 3.74 | 1.63 | 2.3 | 0.02 | 0.3 | 7.98
| Analytical results for Re in OREAS 230b (Certified Value 0.002 ppm) | [18 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.20 | 1.21 | 1.24 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.004 | < 0.01 | 0.003 | 0.008 | 0.002 | 0.002 | 0.003 | 0.003 | 0.002 | 0.0041816 | < 0.05 | 0.003 | < 0.003 | < 0.005 |  |  |  | 1
| 1 | 2 | < 0.002 | < 0.01 | 0.002 | 0.008 | 0.002 | 0.003 | < 0.002 | 0.002 | 0.002 | 0.0051465 | < 0.05 |  | 0.013 | 0.006 |  |  |  | 8
| 1 | 3 | 0.003 | < 0.01 | 0.001 | 0.01 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.0055325 | < 0.05 |  | 0.013 | 0.006 |  |  |  | 16
| 1 | 4 | 0.002 | < 0.01 | 0.003 | 0.009 | 0.002 | 0.003 | 0.002 | 0.003 | 0.002 | 0.0052753 | < 0.05 | 0.003 | < 0.003 | < 0.005 |  |  |  | 0.0024542
| 1 | 5 | 0.002 | < 0.01 | 0.002 | 0.011 | 0.002 | 0.002 | < 0.002 | 0.003 | 0.002 | 0.0045033 | < 0.05 |  | < 0.003 | < 0.005 |  |  |  | 52
| 1 | 6 | 0.003 | < 0.01 | 0.003 | 0.009 | 0.003 | 0.003 | < 0.002 | 0.002 | 0.002 | 0.0045033 | < 0.05 | 0.003 | < 0.003 | 0.006
| Mean |  | 0.0028 |  | 0.0023333 | 0.0091667 | 0.0021667 | 0.0025 | 0.0023333 | 0.0025 | 0.002 | 0.0048571 |  | 0.003 | 0.013 | 0.006
| Median |  | 0.003 |  | 0.0025 | 0.009 | 0.002 | 0.0025 | 0.002 | 0.0025 | 0.002 | 0.0048249 |  | 0.003 | 0.013 | 0.006
| Std Dev. |  | 0.00083666 |  | 0.0008165 | 0.001169 | 0.00040825 | 0.00054772 | 0.00057735 | 0.00054772 | 0 | 0.0005332 |  | 5.3115e-19 | 0 | 1.0623e-18
| Rel.Std.Dev. |  | 29.88% |  | 34.99% | 12.75% | 18.84% | 21.91% | 24.74% | 21.91% | 0% | 10.98% |  | 1.77e-14% | 0% | 1.77e-14%
| PDM3 |  | 14.09% |  | -4.924% | 273.5% | -11.71% | 1.868% | -4.924% | 1.868% | -18.51% | 97.91% |  | 22.24% | 429.7% | 144.5%
| Z-Score (Absolute) |  | 0.18 | 1.66 | 0.31 | 4.72 | 0.43 | 0.18 | 0.8 | 0.18 | 0.55 | 1.55 | 16.37 | 0.18 | 1.9 | 1.1
| Analytical results for S in OREAS 230b (Certified Value 0.272 wt.%) | [18 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 |  | wt.%
|  |  | 4A*OES | AR*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS |  | 3
|  |  |  | 0.25g | [15 empty] | 3
| 1 | 1 | 0.2731 | 0.27808 | 0.26319 | 0.25 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | 0.28 | 0.274 | 0.28 | 0.2778 | 0.239 | 0.16 | 0.28 |  | 1
| 1 | 2 | 0.2752 | 0.27934 | 0.2652 | 0.25 | 0.27 | 0.26 | 0.27 | 0.26 | 0.27 | 0.28 | 0.273 | 0.28 | 0.2778 | 0.2407 | 0.18 | 0.28 |  | 13
| 1 | 3 | 0.2729 | 0.27938 | 0.27077 | 0.26 | 0.28 | 0.26 | 0.27 | 0.27 | 0.27 | 0.28 | 0.27 | 0.27 | 0.277 | 0.2414 | 0.16 | 0.28 |  | 16
| 1 | 4 | 0.2707 | 0.27676 | 0.26156 | 0.26 | 0.27 | 0.26 | 0.27 | 0.27 | 0.27 | 0.28 | 0.273 | 0.28 | 0.2815 | 0.2476 | 0.17 | 0.28 |  | 0.2716
| 1 | 5 | 0.2656 | 0.27659 | 0.26443 | 0.27 | 0.26 | 0.26 | 0.27 | 0.27 | 0.26 | 0.28 | 0.27 | 0.28 | 0.2834 | 0.2398 | 0.17 | 0.28 |  | 53
| 1 | 6 | 0.2694 | 0.27835 | 0.26085 | 0.24 | 0.28 | 0.26 | 0.28 | 0.27 | 0.26 | 0.28 | 0.278 | 0.28 | 0.2793 | 0.2372 | 0.18 | 0.28
| Mean |  | 0.27115 | 0.27808 | 0.26433 | 0.255 | 0.27167 | 0.26167 | 0.27167 | 0.26833 | 0.26667 | 0.28 | 0.273 | 0.27833 | 0.27947 | 0.24095 | 0.17 | 0.28
| Median |  | 0.2718 | 0.27822 | 0.26381 | 0.255 | 0.27 | 0.26 | 0.27 | 0.27 | 0.27 | 0.28 | 0.273 | 0.28 | 0.27855 | 0.24025 | 0.17 | 0.28
| Std Dev. |  | 0.0033863 | 0.0012094 | 0.0035582 | 0.010488 | 0.0075277 | 0.0040825 | 0.0040825 | 0.0040825 | 0.005164 | 0 | 0.0029665 | 0.0040825 | 0.0025009 | 0.0035686 | 0.0089443 | 0
| Rel.Std.Dev. |  | 1.249% | 0.4349% | 1.346% | 4.113% | 2.771% | 1.56% | 1.503% | 1.521% | 1.936% | 0% | 1.087% | 1.467% | 0.8949% | 1.481% | 5.261% | 0%
| PDM3 |  | -0.1648% | 2.389% | -2.675% | -6.111% | 0.02541% | -3.657% | 0.02541% | -1.202% | -1.816% | 3.094% | 0.5163% | 2.48% | 2.897% | -11.28% | -37.41% | 3.094%
| Z-Score (Absolute) |  | 0.02 | 0.64 | 0.68 | 1.58 | 0.02 | 0.94 | 0.02 | 0.3 | 0.46 | 0.83 | 0.15 | 0.67 | 0.78 | 2.93 | 9.77 | 0.83
| Analytical results for Sb in OREAS 230b (Indicative Value 1.44 ppm) | [18 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  |  | 0.25g | [15 empty] | 2
| 1 | 1 | 1.85 | 1.5106 | < 10 | 2.7 | 2.2 | 0.87 | 1.43 | 1.28 | 0.72 | 0.85 | 2.3181 | 1.4 | 1.03 | 0.78 | 9.88 | 3.36 | 1.2 | 1
| 1 | 2 | 1.04 | 1.565 | < 10 | 3.1 | 2.27 | 0.86 | 1.9 | 0.69 | 0.71 | 0.88 | 2.3593 | 1.6 | 1 | 1.13 | 6.74 | 3.4 | 1.1 | 33
| 1 | 3 | 2.03 | 1.4734 | < 10 | 2.4 | 2.23 | 1.27 | 1.52 | 0.65 | 0.74 | 0.89 | 3.1796 | 1.7 | 1 | 0.76 | 9.6 | 3.14 | 1.1 | 16
| 1 | 4 | 3.91 | 1.5063 | < 10 | 2.8 | 2.26 | 0.87 | 1.66 | 1.05 | 0.75 | 0.88 | 3.3031 | 1.5 | 0.98 | 2.52 | 7.11 | 3.06 | 1.4 | 1.4365
| 1 | 5 | 0.56 | 1.4766 | < 10 | 1.8 | 2.25 | 0.89 | 1.49 | 1.07 | 0.75 | 0.87 | 1.9919 | 1.5 | 0.89 | 1.36 | 9.3 | 3.58 | 1.1 | 7
| 1 | 6 | 0.66 | 1.4909 | < 10 | 3 | 2.26 | 0.88 | 1.33 | 0.99 | 0.68 | 0.91 | 2.1824 | 1.4 | 1.02 | 0.77 | 2.28 | 3.04 | 1.1
| Mean |  | 1.675 | 1.5038 |  | 2.6333 | 2.245 | 0.94 | 1.555 | 0.955 | 0.725 | 0.88 | 2.5557 | 1.5167 | 0.98667 | 1.22 | 7.485 | 3.2633 | 1.1667
| Median |  | 1.445 | 1.4986 |  | 2.75 | 2.255 | 0.875 | 1.505 | 1.02 | 0.73 | 0.88 | 2.3387 | 1.5 | 1 | 0.955 | 8.205 | 3.25 | 1.1
| Std Dev. |  | 1.2513 | 0.03354 |  | 0.4761 | 0.025884 | 0.16199 | 0.20087 | 0.24181 | 0.027386 | 0.02 | 0.54774 | 0.1169 | 0.050465 | 0.68197 | 2.8737 | 0.21667 | 0.12111
| Rel.Std.Dev. |  | 74.7% | 2.23% |  | 18.08% | 1.153% | 17.23% | 12.92% | 25.32% | 3.777% | 2.273% | 21.43% | 7.708% | 5.115% | 55.9% | 38.39% | 6.64% | 10.38%
| PDM3 |  | 16.6% | 4.685% |  | 83.32% | 56.28% | -34.56% | 8.25% | -33.52% | -49.53% | -38.74% | 77.92% | 5.581% | -31.31% | -15.07% | 421.1% | 127.2% | -18.78%
| Z-Score (Absolute) |  | 0.19 | 0.02 | 4.07 | 1.31 | 0.85 | 0.67 | 0.04 | 0.66 | 0.93 | 0.74 | 1.22 | 0 | 0.62 | 0.35 | 6.98 | 2.04 | 0.41
| Analytical results for Sc in OREAS 230b (Certified Value 43.1 ppm) | [18 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 42.7 | 43.9991 | 42.6478 | 39 | 43 | 43.5 | 44.9 | 44.2 | 44.4 | 41.3 | 100.6005 | 43 | 41.4 | 42.3 | 37.37 | 50 |  | 1
| 1 | 2 | 42.3 | 44.4526 | 42.4563 | 40 | 42.4 | 42.8 | 45.3 | 44.1 | 44.3 | 42.5 | 138.8418 | 43 | 40.7 | 43.1 | 36.98 | 50.3 |  | 34
| 1 | 3 | 42.4 | 44.5484 | 43.5102 | 44 | 44.5 | 42.5 | 46.6 | 44.9 | 42.3 | 41.7 | 134.224 | 42 | 40.7 | 42.6 | 37.59 | 50.4 |  | 16
| 1 | 4 | 42.4 | 44.6447 | 42.1862 | 42 | 42.9 | 42.7 | 46.9 | 44.5 | 43.7 | 42.4 | 136.9201 | 43 | 41.7 | 43.4 | 37.51 | 50.6 |  | 43.1179
| 1 | 5 | 44.1 | 44.4499 | 42.5513 | 45 | 43.5 | 42.8 | 44.4 | 44.6 | 44.6 | 42.7 | 140.8125 | 43 | 40.9 | 43.1 | 37.76 | 49.7 |  | 54
| 1 | 6 | 42.4 | 44.6509 | 42.282 | 41 | 42.7 | 41.8 | 48.6 | 44.3 | 43.9 | 42.7 | 140.3 | 43 | 41.1 | 43.6 | 37.7 | 49.5
| Mean |  | 42.7167 | 44.4576 | 42.6056 | 41.8333 | 43.1667 | 42.6833 | 46.1167 | 44.4333 | 43.8667 | 42.2167 | 131.9498 | 42.8333 | 41.0833 | 43.0167 | 37.485 | 50.0833
| Median |  | 42.4 | 44.5005 | 42.5038 | 41.5 | 42.95 | 42.75 | 45.95 | 44.4 | 44.1 | 42.45 | 137.8809 | 43 | 41 | 43.1 | 37.55 | 50.15
| Std Dev. |  | 0.69113 | 0.24123 | 0.47454 | 2.3166 | 0.74744 | 0.54924 | 1.5562 | 0.29439 | 0.83586 | 0.58109 | 15.5462 | 0.40825 | 0.40208 | 0.48751 | 0.28346 | 0.42622
| Rel.Std.Dev. |  | 1.618% | 0.5426% | 1.114% | 5.538% | 1.732% | 1.287% | 3.374% | 0.6625% | 1.905% | 1.376% | 11.78% | 0.9531% | 0.9787% | 1.133% | 0.7562% | 0.851%
| PDM3 |  | -0.9306% | 3.107% | -1.188% | -2.979% | 0.113% | -1.008% | 6.955% | 3.051% | 1.736% | -2.09% | 206% | -0.6601% | -4.719% | -0.2349% | -13.06% | 16.15%
| Z-Score (Absolute) |  | 0.14 | 1.02 | 0.21 | 0.72 | 0.16 | 0.16 | 2.12 | 1 | 0.62 | 0.47 | 59.05 | 0.06 | 1.22 | 0.06 | 3.61 | 4.75
| Analytical results for Se in OREAS 230b (Certified Value 0.82 ppm) | [18 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.7 | < 5 | 0.7 | 2.43 | 1 | 1 | 2 | 1 | 1 | 0.78327 | < 1 | < 2 | 0.6 | 1.64 | < 0.5 |  |  | 1
| 1 | 2 | 0.7 | < 5 | 0.6 | 2.54 | 1 | 1 | 2 | 1 | 1 | 0.67247 | < 1 | < 2 |  | 1.64 | < 0.5 |  |  | 0
| 1 | 3 | 0.5 | < 5 | 0.7 | 2.39 | 1 | 1 | 2 | 1 | 1 | 0.64562 | < 1 | < 2 | 0.8 | < 0.05 | < 0.5 |  |  | 16
| 1 | 4 | 0.6 | < 5 | 0.6 | 2.37 | 1 | 1 | 2 | 1 | 1 | 0.55201 | < 1 | < 2 | 0.5 | 0.82 | < 0.5 |  |  | 0.82406
| 1 | 5 | < 0.5 | < 5 | 0.9 | 2.3 | 1 | 1 | 2 | 1 | 1 | 0.64176 | < 1 | < 2 | 0.6 | 0.82 | < 0.5 |  |  | 55
| 1 | 6 | 0.7 | < 5 | 0.5 | 2.53 | 1 | 1 | 2 | 1 | 1 | 0.60393 | < 1 | < 2 | 0.7 | < 0.05 | < 0.5
| Mean |  | 0.64 |  | 0.66667 | 2.4267 | 1 | 1 | 2 | 1 | 1 | 0.64984 |  |  | 0.64 | 1.23
| Median |  | 0.7 |  | 0.65 | 2.41 | 1 | 1 | 2 | 1 | 1 | 0.64369 |  |  | 0.6 | 1.23
| Std Dev. |  | 0.089443 |  | 0.13663 | 0.09395 | 0 | 0 | 0 | 0 | 0 | 0.077561 |  |  | 0.11402 | 0.47343
| Rel.Std.Dev. |  | 13.98% |  | 20.49% | 3.872% | 0% | 0% | 0% | 0% | 0% | 11.94% |  |  | 17.82% | 38.49%
| PDM3 |  | -22.34% |  | -19.1% | 194.5% | 21.35% | 21.35% | 142.7% | 21.35% | 21.35% | -21.14% |  |  | -22.34% | 49.26%
| Z-Score (Absolute) |  | 0.82 | 2.89 | 0.64 | 2.75 | 0 | 0 | 1.93 | 0 | 0 | 0.67 | 0.96 | 0 | 0.69 | 0.33 | 1.44
| Analytical results for Sm in OREAS 230b (Certified Value 2.73 ppm) | [18 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.07; 1.09: 1.09; 1.10: 1.17; 1.11: 1.20 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 2.8221; 1.06: 2.6777; 1.07: 2.9; 1.09: 2.99; 1.10: 2.8174; 1.11: 2.57 |  |  | 1
| 1 | 2 | 1.01: 2.8296; 1.06: 2.5643; 1.07: 2.8; 1.09: 2.71; 1.10: 2.8596; 1.11: 2.75 |  |  | 35
| 1 | 3 | 1.01: 2.8343; 1.06: 2.6214; 1.07: 2.6; 1.09: 2.63; 1.10: 2.9122; 1.11: 2.66 |  |  | 16
| 1 | 4 | 1.01: 2.7653; 1.06: 2.5642; 1.07: 2.4; 1.09: 2.96; 1.10: 2.8412; 1.11: 2.58 |  |  | 2.7332
| 1 | 5 | 1.01: 2.7736; 1.06: 2.6329; 1.07: 2.6; 1.09: 2.87; 1.10: 2.7198; 1.11: 2.63 |  |  | 56
| 1 | 6 | 1.01: 2.7796; 1.06: 2.5085; 1.07: 2.7; 1.09: 3.06; 1.10: 2.8199; 1.11: 2.64
| Mean |  | 1.01: 2.8008; 1.06: 2.5948; 1.07: 2.6667; 1.09: 2.87; 1.10: 2.8283; 1.11: 2.6383
| Median |  | 1.01: 2.8009; 1.06: 2.5929; 1.07: 2.65; 1.09: 2.915; 1.10: 2.8305; 1.11: 2.635
| Std Dev. |  | 1.01: 0.031175; 1.06: 0.060576; 1.07: 0.17512; 1.09: 0.1684; 1.10: 0.063471; 1.11: 0.064936
| Rel.Std.Dev. |  | 1.01: 1.113%; 1.06: 2.334%; 1.07: 6.567%; 1.09: 5.868%; 1.10: 2.244%; 1.11: 2.461%
| PDM3 |  | 1.01: 2.473%; 1.06: -5.061%; 1.07: -2.433%; 1.09: 5.007%; 1.10: 3.483%; 1.11: -3.469%
| Z-Score (Absolute) |  | 1.01: 0.48; 1.06: 0.99; 1.07: 0.48; 1.09: 0.97; 1.10: 0.67; 1.11: 0.68
| Analytical results for Sn in OREAS 230b (Certified Value 0.8 ppm) | [18 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.9 | 0.80689 | < 1 | 2.5 | 0.8 | 0.8 | 0.8 | 1.1 | 0.7 | 0.83783 | 0.8 | 1.5 | 0.9 | 2.55 | 0.8 |  |  | 1
| 1 | 2 | 0.8 | 0.84263 | < 1 | 2.6 | 0.8 | 0.8 | 0.9 | 1.1 | 0.8 | 0.65813 | 0.8 | 1.7 | 1 | 2.79 | 0.8 |  |  | 19
| 1 | 3 | 0.8 | 0.79199 | < 1 | 2.7 | 0.8 | 0.8 | 0.8 | 1 | 0.7 | 0.96055 | 0.7 | 1.5 | 0.9 | 2.32 | 0.9 |  |  | 16
| 1 | 4 | 0.8 | 0.79165 | < 1 | 2.7 | 0.8 | 0.8 | 0.9 | 1.1 | 0.8 | 0.70317 | 0.7 | 1.7 | 0.8 | 3.14 | 0.8 |  |  | 0.80441
| 1 | 5 | 0.8 | 0.77012 | < 1 | 2.7 | 0.7 | 0.7 | 0.8 | 1.1 | 0.8 | 0.92892 | 0.8 | 1.4 | 0.9 | 5.75 | 0.8 |  |  | 57
| 1 | 6 | 0.8 | 0.76022 | < 1 | 2.4 | 0.8 | 0.8 | 0.8 | 1.1 | 0.8 | 0.71229 | 0.7 | 1.4 | 0.9 | 2.72 | 0.8
| Mean |  | 0.81667 | 0.79392 |  | 2.6 | 0.78333 | 0.78333 | 0.83333 | 1.0833 | 0.76667 | 0.80015 | 0.75 | 1.5333 | 0.9 | 3.2117 | 0.81667
| Median |  | 0.8 | 0.79182 |  | 2.65 | 0.8 | 0.8 | 0.8 | 1.1 | 0.8 | 0.77506 | 0.75 | 1.5 | 0.9 | 2.755 | 0.8
| Std Dev. |  | 0.040825 | 0.029162 |  | 0.12649 | 0.040825 | 0.040825 | 0.05164 | 0.040825 | 0.05164 | 0.1273 | 0.054772 | 0.13663 | 0.063246 | 1.2728 | 0.040825
| Rel.Std.Dev. |  | 4.999% | 3.673% |  | 4.865% | 5.212% | 5.212% | 6.197% | 3.768% | 6.736% | 15.91% | 7.303% | 8.91% | 7.027% | 39.63% | 4.999%
| PDM3 |  | 1.524% | -1.304% |  | 223.2% | -2.62% | -2.62% | 3.596% | 34.67% | -4.692% | -0.5292% | -6.764% | 90.62% | 11.88% | 299.3% | 1.524%
| Z-Score (Absolute) |  | 0 | 0.31 | 4.27 | 24.05 | 0.45 | 0.45 | 0.22 | 3.6 | 0.67 | 0.22 | 0.9 | 9.67 | 1.12 | 32.3 | 0
| Analytical results for Sr in OREAS 230b (Certified Value 120 ppm) | [18 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 116.56 | 122.322 | 123.119 | 115 | 117 | 124.5 | 123 | 120.5 | 130.5 | 118 | 163.9071 | 111 | 122.4 | 117 | 109.3 | 119 |  | 1
| 1 | 2 | 115.85 | 120.418 | 121.9607 | 119 | 116 | 121 | 123 | 120 | 132 | 120.5 | 158.3552 | 113 | 129.6 | 118 | 109.1 | 121 |  | 20
| 1 | 3 | 116.55 | 123.246 | 120.9489 | 117 | 116 | 123 | 125 | 122.5 | 128 | 119.5 | 156.845 | 113 | 124.6 | 117 | 108 | 121 |  | 16
| 1 | 4 | 116.46 | 119.957 | 122.0918 | 118 | 116 | 120.5 | 126.5 | 123.5 | 129 | 120 | 158.3585 | 112 | 122.5 | 118 | 110.7 | 121 |  | 119.7834
| 1 | 5 | 118.07 | 122.883 | 120.4796 | 116 | 119 | 120 | 121 | 121.5 | 128.5 | 118.5 | 168.2824 | 115 | 125.6 | 118 | 110.9 | 121 |  | 58
| 1 | 6 | 115.43 | 121.992 | 121.1711 | 120 | 119 | 121 | 129.5 | 120.5 | 127.5 | 122 | 161.6788 | 110 | 125.4 | 120 | 109.4 | 120
| Mean |  | 116.4867 | 121.803 | 121.6285 | 117.5 | 117.1667 | 121.6667 | 124.6667 | 121.4167 | 129.25 | 119.75 | 161.2378 | 112.3333 | 125.0167 | 118 | 109.5667 | 120.5
| Median |  | 116.505 | 122.157 | 121.5659 | 117.5 | 116.5 | 121 | 124 | 121 | 128.75 | 119.75 | 160.0186 | 112.5 | 125 | 118 | 109.35 | 121
| Std Dev. |  | 0.89877 | 1.3326 | 0.95262 | 1.8708 | 1.472 | 1.7224 | 3.0277 | 1.3571 | 1.6956 | 1.4405 | 4.3091 | 1.7512 | 2.64 | 1.0954 | 1.0801 | 0.83666
| Rel.Std.Dev. |  | 0.7716% | 1.094% | 0.7832% | 1.592% | 1.256% | 1.416% | 2.429% | 1.118% | 1.312% | 1.203% | 2.673% | 1.559% | 2.112% | 0.9283% | 0.9858% | 0.6943%
| PDM3 |  | -2.752% | 1.686% | 1.54% | -1.906% | -2.185% | 1.572% | 4.077% | 1.363% | 7.903% | -0.02791% | 34.61% | -6.22% | 4.369% | -1.489% | -8.529% | 0.5982%
| Z-Score (Absolute) |  | 0.84 | 0.16 | 0.13 | 0.65 | 0.71 | 0.13 | 0.7 | 0.09 | 1.56 | 0.23 | 7.58 | 1.62 | 0.76 | 0.56 | 2.14 | 0.09
| Analytical results for Ta in OREAS 230b (Certified Value 0.24 ppm) | [18 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.26 | < 0.5 | 0.3 | 0.38 | 0.23 | 0.22 | 0.25 | 0.24 | 0.21 | 0.48842 | 0.24 | 0.62 | 0.25 | < 0.5 | 0.4 |  |  | 1
| 1 | 2 | 0.25 | < 0.5 | 0.3 | 0.41 | 0.23 | 0.22 | 0.26 | 0.25 | 0.23 | 0.47891 | 0.24 | 0.64 | 0.25 | < 0.5 | 0.3 |  |  | 21
| 1 | 3 | 0.26 | < 0.5 | 0.3 | 0.38 | 0.23 | 0.23 | 0.24 | 0.23 | 0.23 | 0.53785 | 0.23 | 0.66 | 0.23 | < 0.5 | 0.4 |  |  | 16
| 1 | 4 | 0.25 | < 0.5 | 0.3 | 0.39 | 0.22 | 0.23 | 0.25 | 0.24 | 0.22 | 0.71981 | 0.23 | 0.65 | 0.24 | < 0.5 | 0.4 |  |  | 0.23729
| 1 | 5 | 0.25 | < 0.5 | 0.3 | 0.42 | 0.22 | 0.23 | 0.25 | 0.24 | 0.23 | 0.74487 | 0.23 | 0.65 | 0.23 | < 0.5 | 0.3 |  |  | 59
| 1 | 6 | 0.25 | < 0.5 | 0.2 | 0.4 | 0.23 | 0.25 | 0.25 | 0.23 | 0.22 | 0.49824 | 0.24 | 0.63 | 0.25 | < 0.5 | 0.3
| Mean |  | 0.25333 |  | 0.28333 | 0.39667 | 0.22667 | 0.23 | 0.25 | 0.23833 | 0.22333 | 0.57802 | 0.235 | 0.64167 | 0.24167 |  | 0.35
| Median |  | 0.25 |  | 0.3 | 0.395 | 0.23 | 0.23 | 0.25 | 0.24 | 0.225 | 0.51804 | 0.235 | 0.645 | 0.245 |  | 0.35
| Std Dev. |  | 0.005164 |  | 0.040825 | 0.01633 | 0.005164 | 0.010954 | 0.0063246 | 0.0075277 | 0.008165 | 0.12147 | 0.0054772 | 0.01472 | 0.0098319 |  | 0.054772
| Rel.Std.Dev. |  | 2.038% |  | 14.41% | 4.117% | 2.278% | 4.763% | 2.53% | 3.158% | 3.656% | 21.01% | 2.331% | 2.294% | 4.068% |  | 15.65%
| PDM3 |  | 6.76% |  | 19.4% | 67.16% | -4.478% | -3.073% | 5.356% | 0.439% | -5.882% | 143.6% | -0.9658% | 170.4% | 1.844% |  | 47.5%
| Z-Score (Absolute) |  | 0.15 | 0 | NA | 6.59 | 1.05 | 0.9 | 0 | 0.52 | 1.2 | 14.75 | 0.67 | 17.61 | 0.37 | 0 | NA
Results from laboratories 7 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 230b (Certified Value 0.61 ppm) | [18 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.07; 1.09: 1.09; 1.10: 1.17; 1.11: 1.19; 1.12: 1.20; 1.13: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.58634; 1.06: 0.611; 1.07: 0.7; 1.09: 0.7; 1.10: < 0.005; 1.11: 0.62; 1.12: 0.61; 1.13: 0.61 |  |  | 1
| 1 | 2 | 1.01: 0.61043; 1.06: 0.61771; 1.07: 0.7; 1.09: 0.64; 1.10: < 0.005; 1.11: 0.66; 1.12: 0.62; 1.13: 0.62 |  |  | 4
| 1 | 3 | 1.01: 0.59168; 1.06: 0.61413; 1.07: 0.7; 1.09: 0.74; 1.10: < 0.005; 1.11: 0.64; 1.12: 0.6; 1.13: 0.61 |  |  | 16
| 1 | 4 | 1.01: 0.61565; 1.06: 0.58003; 1.07: 0.7; 1.09: 0.71; 1.10: < 0.005; 1.11: 0.63; 1.12: 0.62; 1.13: 0.61 |  |  | 0.61178
| 1 | 5 | 1.01: 0.60364; 1.06: 0.5811; 1.07: 0.7; 1.09: 0.68; 1.10: < 0.005; 1.11: 0.65; 1.12: 0.61; 1.13: 0.62 |  |  | 60
| 1 | 6 | 1.01: 0.58595; 1.06: 0.57564; 1.07: 0.8; 1.09: 0.74; 1.10: < 0.005; 1.11: 0.63; 1.12: 0.61; 1.13: 0.61
| Mean |  | 1.01: 0.59895; 1.06: 0.5966; 1.07: 0.71667; 1.09: 0.70167; 1.11: 0.63833; 1.12: 0.61167; 1.13: 0.61333
| Median |  | 1.01: 0.59766; 1.06: 0.59605; 1.07: 0.7; 1.09: 0.705; 1.11: 0.635; 1.12: 0.61; 1.13: 0.61
| Std Dev. |  | 1.01: 0.012755; 1.06: 0.019565; 1.07: 0.040825; 1.09: 0.038166; 1.11: 0.01472; 1.12: 0.0075277; 1.13: 0.005164
| Rel.Std.Dev. |  | 1.01: 2.13%; 1.06: 3.279%; 1.07: 5.696%; 1.09: 5.439%; 1.11: 2.306%; 1.12: 1.231%; 1.13: 0.842%
| PDM3 |  | 1.01: -2.097%; 1.06: -2.48%; 1.07: 17.15%; 1.09: 14.69%; 1.11: 4.341%; 1.12: -0.01796%; 1.13: 0.2545%
| Z-Score (Absolute) |  | 1.01: 0.57; 1.06: 0.67; 1.07: NA; 1.09: 4.03; 1.10: 27.27; 1.11: 1.19; 1.12: 0; 1.13: 0.07
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 230b (Certified Value 0.12 ppm) | [18 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | < 0.2 | < 1 | < 0.1 | 0.41 | 0.08 | 0.1 | 0.11 | 0.11 | 0.1 | 0.33483 | 0.1 | 0.12 | < 0.03 | < 0.5 | 0.1 |  |  | 1
| 1 | 2 | < 0.2 | < 1 | 0.2 | 0.43 | 0.1 | 0.09 | 0.12 | 0.13 | 0.12 | 0.51054 | 0.1 | 0.13 | < 0.03 | < 0.5 | < 0.1 |  |  | 23
| 1 | 3 | < 0.2 | < 1 | 0.1 | 0.45 | 0.09 | 0.09 | 0.11 | 0.12 | 0.09 | 0.46744 | 0.1 | 0.12 | < 0.03 | < 0.5 | 0.1 |  |  | 16
| 1 | 4 | < 0.2 | < 1 | 0.1 | 0.42 | 0.1 | 0.12 | 0.15 | 0.11 | 0.1 | 0.52048 | 0.1 | 0.14 | < 0.03 | < 0.5 | < 0.1 |  |  | 0.11556
| 1 | 5 | < 0.2 | < 1 | < 0.1 | 0.4 | 0.13 | 0.09 | 0.13 | 0.11 | 0.14 | 0.51718 | 0.1 | 0.14 | 0.71 | < 0.5 | < 0.1 |  |  | 61
| 1 | 6 | < 0.2 | < 1 | < 0.1 | 0.43 | 0.13 | 0.13 | 0.13 | 0.11 | 0.11 | 0.45087 | 0.1 | 0.16 | < 0.03 | < 0.5 | < 0.1
| Mean |  |  |  | 0.13333 | 0.42333 | 0.105 | 0.10333 | 0.125 | 0.115 | 0.11 | 0.46689 | 0.1 | 0.135 | 0.71 |  | 0.1
| Median |  |  |  | 0.1 | 0.425 | 0.1 | 0.095 | 0.125 | 0.11 | 0.105 | 0.48899 | 0.1 | 0.135 | 0.71 |  | 0.1
| Std Dev. |  |  |  | 0.057735 | 0.017512 | 0.020736 | 0.017512 | 0.015166 | 0.0083666 | 0.017889 | 0.070717 | 1.5202e-17 | 0.015166 |  |  | 0
| Rel.Std.Dev. |  |  |  | 43.3% | 4.137% | 19.75% | 16.95% | 12.13% | 7.275% | 16.26% | 15.15% | 1.52e-14% | 11.23% |  |  | 0%
| PDM3 |  |  |  | 15.38% | 266.3% | -9.135% | -10.58% | 8.173% | -0.4808% | -4.808% | 304% | -13.46% | 16.83% | 514.4% |  | -13.46%
| Z-Score (Absolute) |  | 0.79 | 10.56 | NA | 8.39 | 0.65 | 0.7 | 0.08 | 0.37 | 0.51 | 9.62 | NA | 0.2 | 0.08 | 3.47 | NA
Results from laboratories 7, 18 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 230b (Certified Value 0.52 ppm) | [18 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.52 | 0.53723 | 0.53712 | 0.7 | 1.18 | 0.48 | 0.48 | 0.48 | 0.5 | 0.5 | 0.6036 | 0.48 | 0.6 | 0.51 | 0.61 | 0.7 |  | 1
| 1 | 2 | 0.5 | 0.61286 | 0.51969 | 0.6 | 1.15 | 0.48 | 0.51 | 0.45 | 0.51 | 0.51 | 0.57943 | 0.48 | 0.6 | 0.52 | 0.58 | 0.7 |  | 24
| 1 | 3 | 0.52 | 0.58628 | 0.53289 | 1 | 1.08 | 0.5 | 0.48 | 0.45 | 0.47 | 0.51 | 0.57983 | 0.48 | 0.6 | 0.52 | 0.57 | 0.7 |  | 16
| 1 | 4 | 0.5 | 0.60338 | 0.52291 | 0.6 | 1.11 | 0.47 | 0.48 | 0.44 | 0.49 | 0.51 | 0.58605 | 0.49 | 0.8 | 0.5 | 0.6 | 0.6 |  | 0.52039
| 1 | 5 | 0.52 | 0.54198 | 0.55594 | 0.6 | 1.07 | 0.47 | 0.46 | 0.47 | 0.49 | 0.51 | 0.62093 | 0.49 | 0.6 | 0.55 | 0.62 | 0.7 |  | 62
| 1 | 6 | 0.51 | 0.59048 | 0.5302 | 0.6 | 1.07 | 0.5 | 0.51 | 0.45 | 0.49 | 0.52 | 0.60696 | 0.48 | 0.6 | 0.52 | 0.58 | 0.7
| Mean |  | 0.51167 | 0.5787 | 0.53312 | 0.68333 | 1.11 | 0.48333 | 0.48667 | 0.45667 | 0.49167 | 0.51 | 0.59613 | 0.48333 | 0.63333 | 0.52 | 0.59333 | 0.68333
| Median |  | 0.515 | 0.58838 | 0.53155 | 0.6 | 1.095 | 0.48 | 0.48 | 0.45 | 0.49 | 0.51 | 0.59483 | 0.48 | 0.6 | 0.52 | 0.59 | 0.7
| Std Dev. |  | 0.0098319 | 0.031749 | 0.012884 | 0.16021 | 0.046043 | 0.013663 | 0.019664 | 0.015055 | 0.013292 | 0.0063246 | 0.016936 | 0.005164 | 0.08165 | 0.016733 | 0.019664 | 0.040825
| Rel.Std.Dev. |  | 1.922% | 5.486% | 2.417% | 23.45% | 4.148% | 2.827% | 4.041% | 3.297% | 2.703% | 1.24% | 2.841% | 1.068% | 12.89% | 3.218% | 3.314% | 5.974%
| PDM3 |  | -1.675% | 11.21% | 2.448% | 31.31% | 113.3% | -7.12% | -6.48% | -12.24% | -5.519% | -1.996% | 14.56% | -7.12% | 21.7% | -0.07407% | 14.02% | 31.31%
| Z-Score (Absolute) |  | 0 | 1.6 | 0.51 | NA | 14.24 | 0.67 | 0.59 | 1.31 | 0.48 | 0.04 | 2.01 | 0.67 | NA | 0.2 | 1.94 | NA
Results from laboratories 7, 19 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 230b (Certified Value 0.697 wt.%) | [18 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.7059 | 0.70466 | 0.71033 | 0.617 | 0.74 | 0.694 | 0.702 | 0.675 | 0.694 | 0.668 | 0.8907 | 0.71 | 0.67 | 0.6756 | 0.61 | 0.711 |  | 1
| 1 | 2 | 0.7163 | 0.70669 | 0.70037 | 0.64 | 0.73 | 0.68 | 0.697 | 0.66 | 0.699 | 0.692 | 0.91259 | 0.712 | 0.68 | 0.6816 | 0.61 | 0.709 |  | 25
| 1 | 3 | 0.7157 | 0.70779 | 0.72439 | 0.697 | 0.74 | 0.697 | 0.703 | 0.684 | 0.69 | 0.687 | 0.94226 | 0.712 | 0.67 | 0.6664 | 0.61 | 0.714 |  | 16
| 1 | 4 | 0.707 | 0.70557 | 0.70541 | 0.678 | 0.74 | 0.682 | 0.712 | 0.66 | 0.676 | 0.683 | 0.93372 | 0.715 | 0.68 | 0.686 | 0.61 | 0.707 |  | 0.69656
| 1 | 5 | 0.7042 | 0.70795 | 0.70981 | 0.721 | 0.72 | 0.676 | 0.692 | 0.667 | 0.679 | 0.693 | 0.8735 | 0.717 | 0.68 | 0.68 | 0.62 | 0.712 |  | 63
| 1 | 6 | 0.717 | 0.70689 | 0.70585 | 0.658 | 0.75 | 0.681 | 0.73 | 0.668 | 0.676 | 0.69 | 0.89053 | 0.713 | 0.68 | 0.6866 | 0.62 | 0.701
| Mean |  | 0.71102 | 0.70659 | 0.70936 | 0.6685 | 0.73667 | 0.685 | 0.706 | 0.669 | 0.68567 | 0.6855 | 0.90722 | 0.71317 | 0.67667 | 0.67937 | 0.61333 | 0.709
| Median |  | 0.71135 | 0.70679 | 0.70783 | 0.668 | 0.74 | 0.6815 | 0.7025 | 0.6675 | 0.6845 | 0.6885 | 0.90164 | 0.7125 | 0.68 | 0.6808 | 0.61 | 0.71
| Std Dev. |  | 0.0059064 | 0.0012771 | 0.0081946 | 0.038036 | 0.010328 | 0.008438 | 0.01352 | 0.009252 | 0.0099733 | 0.0093113 | 0.027008 | 0.0024833 | 0.005164 | 0.007537 | 0.005164 | 0.0046043
| Rel.Std.Dev. |  | 0.8307% | 0.1807% | 1.155% | 5.69% | 1.402% | 1.232% | 1.915% | 1.383% | 1.455% | 1.358% | 2.977% | 0.3482% | 0.7631% | 1.109% | 0.842% | 0.6494%
| PDM3 |  | 2.076% | 1.441% | 1.838% | -4.028% | 5.758% | -1.659% | 1.356% | -3.956% | -1.564% | -1.587% | 30.24% | 2.384% | -2.856% | -2.468% | -11.95% | 1.786%
| Z-Score (Absolute) |  | 0.65 | 0.46 | 0.58 | 1.16 | 1.74 | 0.46 | 0.43 | 1.14 | 0.43 | 0.44 | 9.01 | 0.74 | 0.82 | 0.7 | 3.52 | 0.56
| Analytical results for Tl in OREAS 230b (Certified Value 0.078 ppm) | [18 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.08 | < 0.1 | 0.08 | 0.1 | 0.07 | 0.08 | 0.08 | 0.09 | 0.07 | 0.18517 | 0.07 | 0.07 | 0.08 | 0.088 | < 0.5 | 0.2 |  | 1
| 1 | 2 | 0.08 | < 0.1 | 0.08 | 0.09 | 0.07 | 0.08 | 0.08 | 0.08 | 0.08 | 0.094576 | 0.07 | 0.08 | 0.08 | 0.061 | < 0.5 | 0.1 |  | 26
| 1 | 3 | 0.08 | < 0.1 | 0.08 | 0.1 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.13918 | 0.07 | 0.08 | 0.08 | 0.062 | < 0.5 | 0.2 |  | 16
| 1 | 4 | 0.09 | < 0.1 | 0.08 | 0.11 | 0.07 | 0.08 | 0.08 | 0.08 | 0.08 | 0.12754 | 0.07 | 0.1 | 0.08 | 0.055 | < 0.5 | 0.2 |  | 0.078315
| 1 | 5 | 0.09 | < 0.1 | 0.08 | 0.09 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.10772 | 0.07 | 0.09 | 0.08 | 0.088 | < 0.5 | 0.2 |  | 64
| 1 | 6 | 0.08 | < 0.1 | 0.08 | 0.1 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.099273 | 0.07 | 0.08 | 0.09 | 0.075 | < 0.5 | 0.1
| Mean |  | 0.083333 |  | 0.08 | 0.098333 | 0.071667 | 0.078333 | 0.08 | 0.081667 | 0.076667 | 0.12558 | 0.07 | 0.083333 | 0.081667 | 0.0715 |  | 0.16667
| Median |  | 0.08 |  | 0.08 | 0.1 | 0.07 | 0.08 | 0.08 | 0.08 | 0.08 | 0.11763 | 0.07 | 0.08 | 0.08 | 0.0685 |  | 0.2
| Std Dev. |  | 0.005164 |  | 0 | 0.0075277 | 0.0040825 | 0.0040825 | 0 | 0.0040825 | 0.005164 | 0.033798 | 0 | 0.010328 | 0.0040825 | 0.014349 |  | 0.05164
| Rel.Std.Dev. |  | 6.197% |  | 0% | 7.655% | 5.696% | 5.212% | 0% | 4.999% | 6.736% | 26.91% | 0% | 12.39% | 4.999% | 20.07% |  | 30.98%
| PDM3 |  | 6.408% |  | 2.151% | 25.56% | -8.489% | 0.02322% | 2.151% | 4.28% | -2.105% | 60.35% | -10.62% | 6.408% | 4.28% | -8.702% |  | 112.8%
| Z-Score (Absolute) |  | 0.67 | 6.07 | 0 | 3.71 | 1.69 | 0.34 | 0 | 0.34 | 0.67 | 9.22 | 2.02 | 0.67 | 0.34 | 1.72 | 34.39 | NA
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 230b (Certified Value 0.38 ppm) | [18 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.06; 1.06: 1.07; 1.07: 1.09; 1.09: 1.17; 1.10: 1.20 |  | ppm
|  |  | 1.01: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 0.35426; 1.06: 0.5; 1.07: 0.41; 1.09: 0.35063; 1.10: 0.38 |  | 1
| 1 | 2 | 1.01: 0.36718; 1.06: 0.4; 1.07: 0.4; 1.09: 0.34493; 1.10: 0.39 |  | 5
| 1 | 3 | 1.01: 0.36968; 1.06: 0.4; 1.07: 0.41; 1.09: 0.36502; 1.10: 0.38 |  | 16
| 1 | 4 | 1.01: 0.3499; 1.06: 0.4; 1.07: 0.41; 1.09: 0.35286; 1.10: 0.39 |  | 0.37867
| 1 | 5 | 1.01: 0.35069; 1.06: 0.4; 1.07: 0.39; 1.09: 0.34839; 1.10: 0.37 |  | 65
| 1 | 6 | 1.01: 0.34538; 1.06: 0.4; 1.07: 0.4; 1.09: 0.33117; 1.10: 0.4
| Mean |  | 1.01: 0.35618; 1.06: 0.41667; 1.07: 0.40333; 1.09: 0.34883; 1.10: 0.385
| Median |  | 1.01: 0.35247; 1.06: 0.4; 1.07: 0.405; 1.09: 0.34951; 1.10: 0.385
| Std Dev. |  | 1.01: 0.0099319; 1.06: 0.040825; 1.07: 0.008165; 1.09: 0.011036; 1.10: 0.010488
| Rel.Std.Dev. |  | 1.01: 2.788%; 1.06: 9.798%; 1.07: 2.024%; 1.09: 3.164%; 1.10: 2.724%
| PDM3 |  | 1.01: -5.939%; 1.06: 10.03%; 1.07: 6.513%; 1.09: -7.879%; 1.10: 1.672%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.06: 0.74; 1.07: 0.43; 1.09: 0.85; 1.10: 0
| Analytical results for U in OREAS 230b (Certified Value 0.18 ppm) | [18 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | AR*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 2
|  |  | 1.06: 0.25g |  | 2
| 1 | 1 | 0.18 | 0.13176 | < 5 | 0.2 | 0.25 | 0.2 | 0.1 | 0.2 | 0.2 | 0.2 | 0.16723 | 0.17 | 0.2 | 0.18 | 0.17 | < 0.5 | 0.2 | 1
| 1 | 2 | 0.18 | 0.12283 | < 5 | 0.2 | 0.26 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16408 | 0.17 | 0.2 | 0.18 | 0.11 | < 0.5 | 0.2 | 28
| 1 | 3 | 0.18 | 0.13491 | < 5 | 0.2 | 0.25 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16341 | 0.17 | 0.2 | 0.17 | 0.07 | < 0.5 | 0.2 | 16
| 1 | 4 | 0.18 | 0.11516 | < 5 | 0.2 | 0.25 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16892 | 0.17 | 0.2 | 0.18 | 0.11 | < 0.5 | 0.2 | 0.17979
| 1 | 5 | 0.17 | 0.13702 | < 5 | 0.1 | 0.23 | 0.2 | 0.1 | 0.2 | 0.2 | 0.2 | 0.14744 | 0.17 | 0.2 | 0.18 | 0.16 | < 0.5 | 0.2 | 66
| 1 | 6 | 0.17 | 0.13972 | < 5 | 0.2 | 0.24 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16532 | 0.17 | 0.2 | 0.18 | 0.18 | < 0.5 | 0.2
| Mean |  | 0.17667 | 0.13023 |  | 0.18333 | 0.24667 | 0.2 | 0.16667 | 0.2 | 0.2 | 0.2 | 0.16273 | 0.17 | 0.2 | 0.17833 | 0.13333 |  | 0.2
| Median |  | 0.18 | 0.13334 |  | 0.2 | 0.25 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1647 | 0.17 | 0.2 | 0.18 | 0.135 |  | 0.2
| Std Dev. |  | 0.005164 | 0.0094068 |  | 0.040825 | 0.010328 | 3.0405e-17 | 0.05164 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 0.0077652 | 0 | 3.0405e-17 | 0.0040825 | 0.043205 |  | 3.0405e-17
| Rel.Std.Dev. |  | 2.923% | 7.223% |  | 22.27% | 4.187% | 1.52e-14% | 30.98% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 4.772% | 0% | 1.52e-14% | 2.289% | 32.4% |  | 1.52e-14%
| PDM3 |  | -1.736% | -27.56% |  | 1.972% | 37.2% | 11.24% | -7.298% | 11.24% | 11.24% | 11.24% | -9.486% | -5.444% | 11.24% | -0.8089% | -25.84% |  | 11.24%
| Z-Score (Absolute) |  | 0.67 | 2.02 | 66.47 | 0.48 | 1.35 | 0 | 0.96 | 0 | 0 | 0 | 1.08 | 0.87 | 0 | 0.63 | 1.93 | 1.44 | 0
| Analytical results for V in OREAS 230b (Certified Value 314 ppm) | [18 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 0
0
| 1 | 1 | 328 | 337.5412 | 309.3094 | 314 | 316 | 316 | 312 | 300 | 324 | 313 | 361.2374 | 305 | 310 | 314 | 321 | 293.5 | 332 | 1
| 1 | 2 | 323 | 333.3332 | 305.1932 | 316 | 317 | 310 | 314 | 300 | 319 | 317 | 374.1689 | 295 | 315 | 320 | 327 | 288.2 | 335 | 29
| 1 | 3 | 329 | 335.2948 | 310.5674 | 310 | 315 | 322 | 313 | 308 | 314 | 310 | 389.4688 | 305 | 313 | 312 | 324 | 290.5 | 336 | 16
| 1 | 4 | 328 | 334.5928 | 307.3318 | 312 | 310 | 308 | 318 | 300 | 309 | 307 | 383.7254 | 300 | 313 | 314 | 326 | 289.8 | 336 | 314.3412
| 1 | 5 | 327 | 333.9124 | 298.3291 | 304 | 319 | 303 | 305 | 297 | 309 | 312 | 355.2236 | 302 | 317 | 318 | 318 | 292.6 | 333 | 67
| 1 | 6 | 335 | 331.18 | 299.4701 | 322 | 320 | 308 | 318 | 306 | 308 | 310 | 366.4731 | 309 | 316 | 314 | 325 | 292.1 | 334
| Mean |  | 328.3333 | 334.3091 | 305.0335 | 313 | 316.1667 | 311.1667 | 313.3333 | 301.8333 | 313.8333 | 311.5 | 371.7162 | 302.6667 | 314 | 315.3333 | 323.5 | 291.1167 | 334.3333
| Median |  | 328 | 334.2526 | 306.2625 | 313 | 316.5 | 309 | 313.5 | 300 | 311.5 | 311 | 370.321 | 303.5 | 314 | 314 | 324.5 | 291.3 | 334.5
| Std Dev. |  | 3.8816 | 2.1166 | 5.1013 | 6.0332 | 3.5449 | 6.7651 | 4.8028 | 4.2151 | 6.4936 | 3.3912 | 13.2264 | 4.8442 | 2.5298 | 3.0111 | 3.3912 | 1.9732 | 1.633
| Rel.Std.Dev. |  | 1.182% | 0.6331% | 1.672% | 1.928% | 1.121% | 2.174% | 1.533% | 1.396% | 2.069% | 1.089% | 3.558% | 1.601% | 0.8057% | 0.9549% | 1.048% | 0.6778% | 0.4884%
| PDM3 |  | 4.451% | 6.352% | -2.961% | -0.4267% | 0.5807% | -1.01% | -0.3206% | -3.979% | -0.1616% | -0.9039% | 18.25% | -3.714% | -0.1085% | 0.3156% | 2.914% | -7.388% | 6.36%
| Z-Score (Absolute) |  | 1.11 | 1.57 | 0.67 | 0.06 | 0.18 | 0.2 | 0.04 | 0.92 | 0 | 0.18 | 4.44 | 0.86 | 0.01 | 0.11 | 0.74 | 1.74 | 1.57
| Analytical results for W in OREAS 230b (Certified Value 9.14 ppm) | [18 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 9 | 9.6767 | 9.1 | 12.2 | 9 | 8.8 | 9.4 | 9.1 | 8.6 | 12.3232 | 9.2 | 9.3 | 9.2 | 9.25 | 9 |  |  | 1
| 1 | 2 | 8.9 | 9.78 | 9.4 | 12.1 | 9.1 | 8.8 | 9.5 | 9.2 | 8.6 | 11.8578 | 9.2 | 9.7 | 9.3 | 8.97 | 8.8 |  |  | 30
| 1 | 3 | 9 | 10.1478 | 9 | 11.3 | 9.2 | 8.8 | 9.4 | 9 | 8.8 | 12.8451 | 9.4 | 9.3 | 9.2 | 9.2 | 9 |  |  | 16
| 1 | 4 | 8.9 | 9.4211 | 8.9 | 12.7 | 8.8 | 8.9 | 9.4 | 9 | 8.7 | 14.5067 | 9.4 | 9.5 | 9.3 | 9.42 | 8.9 |  |  | 9.1405
| 1 | 5 | 9 | 9.9775 | 9.1 | 12 | 8.9 | 8.8 | 9.5 | 9.2 | 8.6 | 15.9707 | 9.5 | 9.4 | 9.2 | 9.05 | 8.6 |  |  | 68
| 1 | 6 | 9.1 | 9.5599 | 9.1 | 11.9 | 9 | 9.3 | 9.3 | 8.9 | 8.4 | 15.7081 | 9.3 | 9.3 | 9.4 | 9.55 | 9
| Mean |  | 8.9833 | 9.7605 | 9.1 | 12.0333 | 9 | 8.9 | 9.4167 | 9.0667 | 8.6167 | 13.8686 | 9.3333 | 9.4167 | 9.2667 | 9.24 | 8.8833
| Median |  | 9 | 9.7283 | 9.1 | 12.05 | 9 | 8.8 | 9.4 | 9.05 | 8.6 | 13.6759 | 9.35 | 9.35 | 9.25 | 9.225 | 8.95
| Std Dev. |  | 0.075277 | 0.26829 | 0.16733 | 0.45461 | 0.14142 | 0.2 | 0.075277 | 0.12111 | 0.13292 | 1.7714 | 0.12111 | 0.16021 | 0.08165 | 0.21872 | 0.16021
| Rel.Std.Dev. |  | 0.838% | 2.749% | 1.839% | 3.778% | 1.571% | 2.247% | 0.7994% | 1.336% | 1.543% | 12.77% | 1.298% | 1.701% | 0.8811% | 2.367% | 1.803%
| PDM3 |  | -1.719% | 6.783% | -0.443% | 31.65% | -1.537% | -2.631% | 3.021% | -0.8077% | -5.731% | 51.73% | 2.11% | 3.021% | 1.38% | 1.089% | -2.813%
| Z-Score (Absolute) |  | 0.72 | 1.46 | 0.39 | 7.85 | 0.67 | 0.96 | 0.5 | 0.49 | 1.75 | 13 | 0.26 | 0.5 | 0.07 | 0 | 1
| Analytical results for Y in OREAS 230b (Certified Value 23.4 ppm) | [18 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 22.24 | 21.9997 | 22.8107 | 24.2 | 22.3 | 23.8 | 23.8 | 22.7 | 24.5 | 23.2 | 25.2505 | 21.3 | 23.6 | 22.22 | 23.82 | 27.7 |  | 1
| 1 | 2 | 21.81 | 23.6928 | 22.547 | 23.6 | 22.4 | 24.3 | 23.3 | 22.6 | 24.3 | 23.7 | 25.3145 | 21.5 | 25 | 22.53 | 23.73 | 28.1 |  | 31
| 1 | 3 | 21.94 | 23.7775 | 22.6021 | 24.2 | 22.1 | 23.9 | 23.8 | 22.7 | 23.7 | 23.3 | 25.9128 | 21.2 | 24.1 | 22.48 | 23.89 | 28.8 |  | 16
| 1 | 4 | 23.09 | 23.6072 | 22.8177 | 24.4 | 22.6 | 23.3 | 23.2 | 23.2 | 25 | 23.8 | 25.6251 | 21.7 | 24.1 | 22.7 | 24.01 | 28.6 |  | 23.3568
| 1 | 5 | 21.81 | 24.0341 | 22.4392 | 24 | 23 | 24 | 23.3 | 23 | 25 | 23.7 | 24.2944 | 21.7 | 24.6 | 22.3 | 23.51 | 27.9 |  | 69
| 1 | 6 | 21.85 | 23.0574 | 22.6547 | 25 | 23 | 23.5 | 24.2 | 22.7 | 23.5 | 24.4 | 24.6909 | 21.3 | 24.2 | 22.58 | 23.5 | 28
| Mean |  | 22.1233 | 23.3614 | 22.6452 | 24.2333 | 22.5667 | 23.8 | 23.6 | 22.8167 | 24.3333 | 23.6833 | 25.1814 | 21.45 | 24.2667 | 22.4683 | 23.7433 | 28.1833
| Median |  | 21.895 | 23.65 | 22.6284 | 24.2 | 22.5 | 23.85 | 23.55 | 22.7 | 24.4 | 23.7 | 25.2825 | 21.4 | 24.15 | 22.505 | 23.775 | 28.05
| Std Dev. |  | 0.50055 | 0.7406 | 0.14911 | 0.46332 | 0.37238 | 0.35777 | 0.39497 | 0.23166 | 0.63456 | 0.42622 | 0.59695 | 0.21679 | 0.48028 | 0.17893 | 0.20607 | 0.42622
| Rel.Std.Dev. |  | 2.263% | 3.17% | 0.6585% | 1.912% | 1.65% | 1.503% | 1.674% | 1.015% | 2.608% | 1.8% | 2.371% | 1.011% | 1.979% | 0.7964% | 0.8679% | 1.512%
| PDM3 |  | -5.281% | 0.01986% | -3.046% | 3.753% | -3.383% | 1.898% | 1.041% | -2.313% | 4.181% | 1.398% | 7.812% | -8.164% | 3.895% | -3.804% | 1.655% | 20.66%
| Z-Score (Absolute) |  | 1.35 | 0.25 | 0.89 | 0.53 | 0.96 | 0.14 | 0.04 | 0.73 | 0.62 | 0.04 | 1.37 | 1.95 | 0.56 | 1.04 | 0.09 | 4.04
| Analytical results for Yb in OREAS 230b (Certified Value 2.53 ppm) | [18 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.07; 1.09: 1.09; 1.10: 1.17; 1.11: 1.19; 1.12: 1.20; 1.13: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 2.6906; 1.06: 2.3272; 1.07: 2.4; 1.09: 2.54; 1.10: 2.1298; 1.11: 2.5; 1.12: 2.39; 1.13: 2.46 |  | 1
| 1 | 2 | 1.01: 2.6875; 1.06: 2.4127; 1.07: 2.5; 1.09: 2.71; 1.10: 2.0869; 1.11: 2.6; 1.12: 2.48; 1.13: 2.43 |  | 6
| 1 | 3 | 1.01: 2.6815; 1.06: 2.3904; 1.07: 2.4; 1.09: 2.67; 1.10: 2.1764; 1.11: 2.5; 1.12: 2.43; 1.13: 2.46 |  | 16
| 1 | 4 | 1.01: 2.7156; 1.06: 2.2765; 1.07: 2.6; 1.09: 2.89; 1.10: 2.1592; 1.11: 2.4; 1.12: 2.43; 1.13: 2.47 |  | 2.5266
| 1 | 5 | 1.01: 2.705; 1.06: 2.4744; 1.07: 2.3; 1.09: 2.72; 1.10: 2.0312; 1.11: 2.6; 1.12: 2.41; 1.13: 2.45 |  | 70
| 1 | 6 | 1.01: 2.6827; 1.06: 2.3426; 1.07: 2.8; 1.09: 2.65; 1.10: 2.0229; 1.11: 2.5; 1.12: 2.56; 1.13: 2.48
| Mean |  | 1.01: 2.6938; 1.06: 2.3706; 1.07: 2.5; 1.09: 2.6967; 1.10: 2.1011; 1.11: 2.5167; 1.12: 2.45; 1.13: 2.4583
| Median |  | 1.01: 2.689; 1.06: 2.3665; 1.07: 2.45; 1.09: 2.69; 1.10: 2.1084; 1.11: 2.5; 1.12: 2.43; 1.13: 2.46
| Std Dev. |  | 1.01: 0.013585; 1.06: 0.069951; 1.07: 0.17889; 1.09: 0.11448; 1.10: 0.06491; 1.11: 0.075277; 1.12: 0.061644; 1.13: 0.017224
| Rel.Std.Dev. |  | 1.01: 0.5043%; 1.06: 2.951%; 1.07: 7.155%; 1.09: 4.245%; 1.10: 3.089%; 1.11: 2.991%; 1.12: 2.516%; 1.13: 0.7006%
| PDM3 |  | 1.01: 6.619%; 1.06: -6.172%; 1.07: -1.052%; 1.09: 6.731%; 1.10: -16.84%; 1.11: -0.3927%; 1.12: -3.031%; 1.13: -2.702%
| Z-Score (Absolute) |  | 1.01: 1.98; 1.06: 1; 1.07: 0.19; 1.09: 2.01; 1.10: 3.49; 1.11: 0.35; 1.12: 0.27; 1.13: 0.19
| Analytical results for Zn in OREAS 230b (Certified Value 109 ppm) | [18 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 0
0
| 1 | 1 | 117 | 100.0072 | 109.306 | 98 | 105 | 115 | 118 | 111 | 117 | 116 | 153.1333 | 108 | 106 | 105 | 98 | 102.8 | 107 | 1
| 1 | 2 | 115 | 101.7856 | 107.3179 | 102 | 104 | 113 | 118 | 110 | 117 | 117 | 145.5331 | 108 | 105 | 106 | 98 | 107.1 | 107 | 33
| 1 | 3 | 114 | 100.6832 | 110.4449 | 110 | 105 | 117 | 117 | 112 | 117 | 117 | 146.7524 | 106 | 104 | 105 | 98 | 110.1 | 108 | 16
| 1 | 4 | 115 | 101.3656 | 109.552 | 109 | 104 | 113 | 118 | 111 | 114 | 117 | 151.2391 | 108 | 103 | 106 | 99 | 110.8 | 109 | 109.097
| 1 | 5 | 122 | 102.7936 | 107.2188 | 114 | 108 | 111 | 116 | 111 | 114 | 117 | 156.8852 | 108 | 106 | 106 | 97 | 102.7 | 108 | 71
| 1 | 6 | 116 | 101.5984 | 106.9376 | 104 | 111 | 113 | 122 | 110 | 113 | 117 | 148.1837 | 109 | 108 | 107 | 98 | 109.8 | 108
| Mean |  | 116.5 | 101.3723 | 108.4629 | 106.1667 | 106.1667 | 113.6667 | 118.1667 | 110.8333 | 115.3333 | 116.8333 | 150.2878 | 107.8333 | 105.3333 | 105.8333 | 98 | 107.2167 | 107.8333
| Median |  | 115.5 | 101.482 | 108.312 | 106.5 | 105 | 113 | 118 | 111 | 115.5 | 117 | 149.7114 | 108 | 105.5 | 106 | 98 | 108.45 | 108
| Std Dev. |  | 2.881 | 0.95712 | 1.4839 | 5.8793 | 2.7869 | 2.0656 | 2.0412 | 0.75277 | 1.8619 | 0.40825 | 4.2909 | 0.98319 | 1.7512 | 0.75277 | 0.63246 | 3.681 | 0.75277
| Rel.Std.Dev. |  | 2.473% | 0.9442% | 1.368% | 5.538% | 2.625% | 1.817% | 1.727% | 0.6792% | 1.614% | 0.3494% | 2.855% | 0.9118% | 1.663% | 0.7113% | 0.6454% | 3.433% | 0.6981%
| PDM3 |  | 6.786% | -7.081% | -0.5812% | -2.686% | -2.686% | 4.189% | 8.313% | 1.592% | 5.716% | 7.091% | 37.76% | -1.158% | -3.45% | -2.992% | -10.17% | -1.724% | -1.158%
| Z-Score (Absolute) |  | 1.95 | 1.45 | 0.14 | 0.37 | 0.37 | 1.31 | 2.32 | 0.67 | 1.69 | 2.02 | 9.54 | 0 | 0.56 | 0.45 | 2.21 | 0.14 | 0
| Analytical results for Zr in OREAS 230b (Certified Value 56 ppm) | [18 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.24 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 58.2 | 45.0605 | 57.6547 | 54 | 58 | 56.1 | 62.6 | 53.2 | 56.5 | 55.4 | 74.7466 | 60 | 83.6 | 58.2 | 76.8 | 73.5 |  | 1
| 1 | 2 | 59.4 | 44.378 | 55.9895 | 58 | 57 | 62 | 64.6 | 53 | 54.8 | 52.1 | 70.7904 | 62.9 | 85.5 | 54.2 | 70.84 | 73.2 |  | 34
| 1 | 3 | 55.1 | 45.7325 | 57.3137 | 54 | 64 | 55.1 | 58.1 | 53.2 | 56.7 | 49.8 | 72.0653 | 59 | 82.8 | 55.6 | 71.41 | 73.2 |  | 16
| 1 | 4 | 53.6 | 45.491 | 56.5232 | 56 | 64 | 55.2 | 57.7 | 55 | 60.3 | 52.3 | 70.227 | 61.5 | 83.1 | 54.9 | 77.27 | 72.3 |  | 55.8758
| 1 | 5 | 54.4 | 45.239 | 54.6489 | 56 | 56 | 57.4 | 59 | 53.4 | 56.8 | 55.1 | 75.4065 | 59.3 | 83.3 | 53.4 | 76.55 | 71.8 |  | 72
| 1 | 6 | 55.5 | 45.4175 | 55.3069 | 58 | 60 | 58.7 | 61 | 54.3 | 52.4 | 53.1 | 71.3559 | 58.8 | 84.9 | 60.4 | 68.22 | 71.6
| Mean |  | 56.0333 | 45.2197 | 56.2395 | 56 | 59.8333 | 57.4167 | 60.5 | 53.6833 | 56.25 | 52.9667 | 72.4319 | 60.25 | 83.8667 | 56.1167 | 73.515 | 72.6
| Median |  | 55.3 | 45.3282 | 56.2564 | 56 | 59 | 56.75 | 60 | 53.3 | 56.6 | 52.7 | 71.7106 | 59.65 | 83.45 | 55.25 | 73.98 | 72.75
| Std Dev. |  | 2.2704 | 0.47114 | 1.1575 | 1.7889 | 3.4881 | 2.6332 | 2.732 | 0.79099 | 2.6037 | 2.0839 | 2.1473 | 1.6257 | 1.082 | 2.6641 | 3.84 | 0.80747
| Rel.Std.Dev. |  | 4.052% | 1.042% | 2.058% | 3.194% | 5.83% | 4.586% | 4.516% | 1.473% | 4.629% | 3.934% | 2.965% | 2.698% | 1.29% | 4.748% | 5.223% | 1.112%
| PDM3 |  | 0.282% | -19.07% | 0.6509% | 0.2223% | 7.083% | 2.758% | 8.276% | -3.924% | 0.6698% | -5.206% | 29.63% | 7.828% | 50.09% | 0.4311% | 31.57% | 29.93%
| Z-Score (Absolute) |  | 0.16 | 2.39 | 0.12 | 0.17 | 0.62 | 0.12 | 0.75 | 0.65 | 0.12 | 0.79 | 3.2 | 0.7 | 5.55 | 0.15 | 3.43 | 3.24
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C513: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C714: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C952: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 230b (Certified Value 0.181 ppm) | [16 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.167 | 0.19453 | 0.17 | 0.18 | 0.18 | 0.32 | 0.19 | 0.18 | 0.39808 | 0.14 | 0.18 | < 0.3 | 1.11 | 0.3 | 0.19 | 1
| 1 | 2 | 0.214 | 0.20282 | 0.165 | 0.17 | 0.18 | 0.18 | 0.24 | 0.18 | 0.39263 | 0.15 | 0.18 | < 0.3 | 1.07 | 0.2 | 0.18 | 19
| 1 | 3 | 0.173 | 0.19943 | 0.166 | 0.17 | 0.18 | 0.2 | 0.19 | 0.19 | 0.50263 | 0.15 | 0.17 | < 0.3 | 1.12 | 0.2 | 0.19 | 16
| 1 | 4 | 0.174 | 0.19617 | 0.177 | 0.17 | 0.18 | 0.24 | 0.18 | 0.19 | 0.55021 | 0.16 | 0.18 | < 0.3 | 1.15 | 0.2 | 0.28 | 0.18068
| 1 | 5 | 0.166 | 0.18725 | 0.178 | 0.18 | 0.18 | 0.22 | 0.18 | 0.19 | 0.45025 | 0.15 | 0.19 | 0.4 | 1.02 | 0.2 | 0.21 | 74
| 1 | 6 |  | 0.20346 | 0.174 | 0.19 | 0.18 | 0.21 | 0.18 | 0.18 | 0.41678 | 0.14 | 0.18 | < 0.3 | 1.03 | 0.2 | 0.19
| Mean |  | 0.1788 | 0.19728 | 0.17167 | 0.17667 | 0.18 | 0.22833 | 0.19333 | 0.185 | 0.45176 | 0.14833 | 0.18 | 0.4 | 1.0833 | 0.21667 | 0.20667
| Median |  | 0.173 | 0.1978 | 0.172 | 0.175 | 0.18 | 0.215 | 0.185 | 0.185 | 0.43352 | 0.15 | 0.18 | 0.4 | 1.09 | 0.2 | 0.19
| Std Dev. |  | 0.019992 | 0.0060456 | 0.0055377 | 0.008165 | 3.0405e-17 | 0.04916 | 0.023381 | 0.0054772 | 0.06306 | 0.0075277 | 0.0063246 |  | 0.052026 | 0.040825 | 0.037238
| Rel.Std.Dev. |  | 11.18% | 3.065% | 3.226% | 4.622% | 1.689e-14% | 21.53% | 12.09% | 2.961% | 13.96% | 5.075% | 3.514% |  | 4.802% | 18.84% | 18.02%
| PDM3 |  | -1.039% | 9.188% | -4.987% | -2.219% | -0.3745% | 26.38% | 7.005% | 2.393% | 150% | -17.9% | -0.3745% | 121.4% | 499.6% | 19.92% | 14.38%
| Z-Score (Absolute) |  | 0.61 | 0.57 | 1.06 | 0.74 | 0.53 | 2.54 | 0.32 | 0.21 | 16.76 | 2.54 | 0.53 | 0.21 | 56.93 | NA | 1.17
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 230b (Certified Value 3.4 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 3.2829 | 3.4258 | 3.41 | 3.43 | 3.42 | 3.22 | 3.55 | 3.19 | 4.3956 | 3.82 | 3.33 | 3.8 | 3.11 | 3.16 |  | 1
| 1 | 2 | 3.1593 | 3.4327 | 3.37 | 3.44 | 3.45 | 3.27 | 3.57 | 3.09 | 4.4572 | 3.8 | 3.34 | 3.88 | 3.16 | 3.08 |  | 0
| 1 | 3 | 3.3105 | 3.4748 | 3.38 | 3.43 | 3.42 | 3.43 | 3.57 | 3.31 | 4.1387 | 3.85 | 3.32 | 3.7 | 3.11 | 3.09 |  | 16
| 1 | 4 | 3.3273 | 3.4114 | 3.38 | 3.36 | 3.35 | 3.35 | 3.59 | 3.36 | 4.2555 | 3.8 | 3.35 | 3.82 | 3.16 | 3.09 |  | 3.3989
| 1 | 5 | 3.3572 | 3.3317 | 3.47 | 3.43 | 3.35 | 3.39 | 3.52 | 3.12 | 4.0878 | 3.84 | 3.37 | 3.79 | 3.16 | 3.13 |  | 75
| 1 | 6 | 3.0997 | 3.4488 | 3.36 | 3.41 | 3.39 | 3.36 | 3.62 | 3.06 | 4.2236 | 3.85 | 3.33 | 3.76 | 3.12 | 3.11
| Mean |  | 3.2562 | 3.4209 | 3.395 | 3.4167 | 3.3967 | 3.3367 | 3.57 | 3.1883 | 4.2598 | 3.8267 | 3.34 | 3.7917 | 3.1367 | 3.11
| Median |  | 3.2967 | 3.4293 | 3.38 | 3.43 | 3.405 | 3.355 | 3.57 | 3.155 | 4.2396 | 3.83 | 3.335 | 3.795 | 3.14 | 3.1
| Std Dev. |  | 0.10275 | 0.048769 | 0.040373 | 0.029439 | 0.040825 | 0.077889 | 0.034059 | 0.12254 | 0.14353 | 0.023381 | 0.017889 | 0.060139 | 0.02582 | 0.030332
| Rel.Std.Dev. |  | 3.156% | 1.426% | 1.189% | 0.8616% | 1.202% | 2.334% | 0.954% | 3.843% | 3.369% | 0.611% | 0.5356% | 1.586% | 0.8232% | 0.9753%
| PDM3 |  | -4.199% | 0.6473% | -0.114% | 0.5235% | -0.06495% | -1.83% | 5.035% | -6.194% | 25.33% | 12.59% | -1.732% | 11.56% | -7.715% | -8.499%
| Z-Score (Absolute) |  | 0.6 | 0.11 | 0 | 0.09 | 0 | 0.25 | 0.75 | 0.89 | 3.71 | 1.85 | 0.24 | 1.7 | 1.11 | 1.23
| Analytical results for As in OREAS 230b (Certified Value 25.7 ppm) | [16 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 25.137 | 27.3731 | 26.1 | 25.3 | 26.9 | 26.8 | 27.9 | 25 | 25.0559 | 25 | 27 | 25.7 | 6.5 | 24.6 | 26.3 | 1
| 1 | 2 | 25.672 | 27.0642 | 25.4 | 25.5 | 27.5 | 26.9 | 28.2 | 24.3 | 24.5401 | 25 | 24 | 23.4 | 6.8 | 24.7 | 25.6 | 20
| 1 | 3 | 25.464 | 25.7719 | 25.5 | 24.7 | 26.9 | 25.8 | 29 | 25.2 | 23.5368 | 25 | 26 | 25 | 6.9 | 24.6 | 26 | 16
| 1 | 4 | 25.379 | 27.2021 | 25.9 | 25.1 | 27.1 | 27.1 | 27.7 | 25.2 | 22.1094 | 25 | 26 | 23.8 | 6 | 24.5 | 26.6 | 25.6842
| 1 | 5 | 27.357 | 26.5488 | 27 | 25.7 | 27 | 27.2 | 28.6 | 22.8 | 22.9455 | 25 | 27 | 24.9 | 6.7 | 24.1 | 26.9 | 76
| 1 | 6 | 24 | 27.1964 | 25.7 | 25.3 | 27.2 | 26.7 | 28 | 24.8 | 23.3169 | 25 | 24 | 23.6 | 6.5 | 24.5 | 26
| Mean |  | 25.5015 | 26.8594 | 25.9333 | 25.2667 | 27.1 | 26.75 | 28.2333 | 24.55 | 23.5841 | 25 | 25.6667 | 24.4 | 6.5667 | 24.5 | 26.2333
| Median |  | 25.4215 | 27.1303 | 25.8 | 25.3 | 27.05 | 26.85 | 28.1 | 24.9 | 23.4269 | 25 | 26 | 24.35 | 6.6 | 24.55 | 26.15
| Std Dev. |  | 1.084 | 0.60265 | 0.58195 | 0.34448 | 0.22804 | 0.501 | 0.48442 | 0.92033 | 1.0711 | 0 | 1.3663 | 0.92736 | 0.32042 | 0.20976 | 0.46762
| Rel.Std.Dev. |  | 4.251% | 2.244% | 2.244% | 1.363% | 0.8415% | 1.873% | 1.716% | 3.749% | 4.542% | 0% | 5.323% | 3.801% | 4.879% | 0.8562% | 1.783%
| PDM3 |  | -0.7112% | 4.576% | 0.9701% | -1.626% | 5.512% | 4.15% | 9.925% | -4.416% | -8.176% | -2.664% | -0.06815% | -5% | -74.43% | -4.611% | 2.138%
| Z-Score (Absolute) |  | 0.05 | 0.81 | 0.22 | 0.2 | 0.97 | 0.74 | 1.69 | 0.66 | 1.27 | 0.37 | 0.05 | 0.75 | NA | 0.69 | 0.41
| Analytical results for B in OREAS 230b (Indicative Value 13.5 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 0.5g |  |  |  |  | 1
| 1 | 1 | 102.471 | 14.7787 | 26 | 10 | 10 | 10 | 20 | 10 | 13.3303 | 17 | 9 |  |  |  |  | 1
| 1 | 2 | 135.326 | 14.6302 | 22 | 10 | 10 | 20 | 10 | 10 | 9.7732 | 17 | 9 |  |  |  |  | 3
| 1 | 3 | 110.815 | 16.8249 | 20 | 10 | 10 | < 10 | 20 | 10 | 11.4887 | 16 | 9 |  |  |  |  | 16
| 1 | 4 | 106.695 | 14.399 | 24 | 10 | 10 | 10 | 20 | 10 | 11.144 | 16 | 9 |  |  |  |  | 13.5038
| 1 | 5 | 119.791 | 14.5672 | 23 | 10 | 10 | 10 | 20 | 10 | 10.8682 | 16 | 9 |  |  |  |  | 9
| 1 | 6 | 121 | 16.9356 | 22 | 10 | 10 | 10 | 20 | 10 | 11.4879 | 15 | 9
| Mean |  | 116.0163 | 15.3559 | 22.8333 | 10 | 10 | 12 | 18.3333 | 10 | 11.3487 | 16.1667 | 9
| Median |  | 115.303 | 14.7045 | 22.5 | 10 | 10 | 10 | 20 | 10 | 11.3159 | 16 | 9
| Std Dev. |  | 11.911 | 1.1875 | 2.0412 | 0 | 0 | 4.4721 | 4.0825 | 0 | 1.1594 | 0.75277 | 0
| Rel.Std.Dev. |  | 10.27% | 7.733% | 8.94% | 0% | 0% | 37.27% | 22.27% | 0% | 10.22% | 4.656% | 0%
| PDM3 |  | 759.1% | 13.72% | 69.09% | -25.95% | -25.95% | -11.14% | 35.76% | -25.95% | -15.96% | 19.72% | -33.35%
| Z-Score (Absolute) |  | 30.05 | 1.15 | 3.3 | 0.39 | 0.39 | 0.15 | 2.01 | 0.39 | 0 | 1.38 | 0.67
| Analytical results for Ba in OREAS 230b (Certified Value 19.7 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 17.841 | 18.3833 | 20 | 20 | 20 | 20 | 20 | 20 | 17.5409 | 20 | 19 | 19.9 | 16 | 14 |  | 1
| 1 | 2 | 17.772 | 18.5298 | 20.6 | 20 | 20 | 20 | 20 | 20 | 17.379 | 20 | 18 | 20.8 | 16 | 14 |  | 21
| 1 | 3 | 17.174 | 18.3182 | 19.4 | 20 | 20 | 20 | 20 | 20 | 20.8315 | 20 | 18 | 20.2 | 14 | 15 |  | 16
| 1 | 4 | 17.669 | 18.8303 | 20 | 20 | 20 | 20 | 20 | 20 | 18.6198 | 20 | 19 | 19.9 | 19 | 14 |  | 19.7329
| 1 | 5 | 17.805 | 18.324 | 20.2 | 20 | 20 | 20 | 20 | 20 | 17.1313 | 20 | 18 | 20.6 | 17 | 14 |  | 77
| 1 | 6 | 17 | 18.4882 | 21 | 20 | 20 | 20 | 20 | 20 | 17.9185 | 20 | 18 | 20.5 | 17 | 14
| Mean |  | 17.5435 | 18.479 | 20.2 | 20 | 20 | 20 | 20 | 20 | 18.2368 | 20 | 18.3333 | 20.3167 | 16.5 | 14.1667
| Median |  | 17.7205 | 18.4358 | 20.1 | 20 | 20 | 20 | 20 | 20 | 17.7297 | 20 | 18 | 20.35 | 16.5 | 14
| Std Dev. |  | 0.36243 | 0.1924 | 0.55136 | 0 | 0 | 0 | 0 | 0 | 1.3728 | 0 | 0.5164 | 0.37639 | 1.6432 | 0.40825
| Rel.Std.Dev. |  | 2.066% | 1.041% | 2.73% | 0% | 0% | 0% | 0% | 0% | 7.527% | 0% | 2.817% | 1.853% | 9.959% | 2.882%
| PDM3 |  | -11.1% | -6.354% | 2.367% | 1.354% | 1.354% | 1.354% | 1.354% | 1.354% | -7.582% | 1.354% | -7.093% | 2.958% | -16.38% | -28.21%
| Z-Score (Absolute) |  | 6.41 | 3.97 | 0.52 | 0 | 0 | 0 | 0 | 0 | 4.6 | 0 | 4.35 | 0.83 | 9.14 | 15.23
| Analytical results for Be in OREAS 230b (Certified Value 0.15 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 0.08 | 0.21484 | 0.2 | 0.15 | 0.15 | 0.17 | 0.15 | 0.17 | < 0.5 | 0.2 | < 0.5 | < 1 | 0.12 |  |  | 1
| 1 | 2 | 0.102 | 0.20793 | 0.2 | 0.15 | 0.15 | 0.17 | 0.15 | 0.17 | < 0.5 | 0.2 | < 0.5 | < 1 | 0.12 |  |  | 22
| 1 | 3 | 0.091 | 0.20227 | 0.3 | 0.15 | 0.15 | 0.15 | 0.15 | 0.18 | < 0.5 | 0.2 | < 0.5 | < 1 | 0.14 |  |  | 16
| 1 | 4 | 0.088 | 0.21192 | 0.2 | 0.14 | 0.15 | 0.18 | 0.15 | 0.17 | < 0.5 | 0.2 | < 0.5 | < 1 | 0.19 |  |  | 0.15086
| 1 | 5 | 0.089 | 0.20567 | 0.2 | 0.15 | 0.15 | 0.17 | 0.14 | 0.17 | < 0.5 | 0.2 | < 0.5 | < 1 | 0.11 |  |  | 78
| 1 | 6 |  | 0.20883 | 0.2 | 0.15 | 0.14 | 0.18 | 0.15 | 0.18 | < 0.5 | 0.2 | < 0.5 | < 1 | 0.11
| Mean |  | 0.09 | 0.20858 | 0.21667 | 0.14833 | 0.14833 | 0.17 | 0.14833 | 0.17333 |  | 0.2 |  |  | 0.13167
| Median |  | 0.089 | 0.20838 | 0.2 | 0.15 | 0.15 | 0.17 | 0.15 | 0.17 |  | 0.2 |  |  | 0.12
| Std Dev. |  | 0.0079057 | 0.0044523 | 0.040825 | 0.0040825 | 0.0040825 | 0.010954 | 0.0040825 | 0.005164 |  | 3.0405e-17 |  |  | 0.030605
| Rel.Std.Dev. |  | 8.784% | 2.135% | 18.84% | 2.752% | 2.752% | 6.444% | 2.752% | 2.979% |  | 1.52e-14% |  |  | 23.24%
| PDM3 |  | -40.34% | 38.25% | 43.62% | -1.677% | -1.677% | 12.68% | -1.677% | 14.89% |  | 32.57% |  |  | -12.72%
| Z-Score (Absolute) |  | 1.41 | 0.68 | NA | 0.38 | 0.38 | 0 | 0.38 | 0.06 | 1.41 | NA | 1.41 | 5.8 | 0.67
Results from laboratories 7 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 230b (Certified Value 0.24 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.233 | 0.23861 | 0.25 | 0.21 | 0.23 | 0.25 | 0.23 | 0.25 | 0.24225 | 0.23 | 0.24 | 0.26 | < 0.5 | 0.2 | 0.27 | 1
| 1 | 2 | 0.245 | 0.25824 | 0.24 | 0.22 | 0.23 | 0.23 | 0.24 | 0.24 | 0.23507 | 0.23 | 0.24 | 0.23 | < 0.5 | 0.2 | 0.27 | 23
| 1 | 3 | 0.222 | 0.24346 | 0.24 | 0.22 | 0.23 | 0.23 | 0.23 | 0.25 | 0.24175 | 0.23 | 0.23 | 0.24 | < 0.5 | 0.2 | 0.27 | 16
| 1 | 4 | 0.234 | 0.24334 | 0.26 | 0.22 | 0.23 | 0.23 | 0.23 | 0.24 | 0.24255 | 0.23 | 0.24 | 0.23 | < 0.5 | 0.2 | 0.28 | 0.23913
| 1 | 5 | 0.224 | 0.23577 | 0.26 | 0.22 | 0.22 | 0.25 | 0.23 | 0.25 | 0.23905 | 0.23 | 0.24 | 0.26 | < 0.5 | 0.2 | 0.26 | 79
| 1 | 6 |  | 0.24377 | 0.25 | 0.21 | 0.24 | 0.25 | 0.24 | 0.25 | 0.25163 | 0.23 | 0.23 | 0.32 | < 0.5 | 0.2 | 0.25
| Mean |  | 0.2316 | 0.24387 | 0.25 | 0.21667 | 0.23 | 0.24 | 0.23333 | 0.24667 | 0.24205 | 0.23 | 0.23667 | 0.25667 |  | 0.2 | 0.26667
| Median |  | 0.233 | 0.2434 | 0.25 | 0.22 | 0.23 | 0.24 | 0.23 | 0.25 | 0.242 | 0.23 | 0.24 | 0.25 |  | 0.2 | 0.27
| Std Dev. |  | 0.0091815 | 0.0077457 | 0.0089443 | 0.005164 | 0.0063246 | 0.010954 | 0.005164 | 0.005164 | 0.005476 | 0 | 0.005164 | 0.033862 |  | 3.0405e-17 | 0.010328
| Rel.Std.Dev. |  | 3.964% | 3.176% | 3.578% | 2.383% | 2.75% | 4.564% | 2.213% | 2.094% | 2.262% | 0% | 2.182% | 13.19% |  | 1.52e-14% | 3.873%
| PDM3 |  | -3.147% | 1.982% | 4.547% | -9.392% | -3.817% | 0.3653% | -2.423% | 3.153% | 1.223% | -3.817% | -1.029% | 7.335% |  | -16.36% | 11.52%
| Z-Score (Absolute) |  | 0.71 | 0.21 | 0.67 | 1.83 | 0.83 | 0.08 | 0.58 | 0.42 | 0.08 | 0.83 | 0.33 | 1.18 | 0.67 | NA | 1.93
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 230b (Certified Value 3.11 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 2.5094 | 3.4058 | 4.36 | 3.14 | 3.02 | 3.12 | 3.41 | 3.46 | 3.1608 | 3.77 | 2.92 | 2.8469 | 3.23 | 2.99 | 2.7 | 1
| 1 | 2 | 2.5314 | 3.3996 | 3.93 | 3.17 | 3.06 | 3.18 | 3.42 | 3.35 | 3.2015 | 3.78 | 2.96 | 2.7579 | 3.26 | 3.04 | 2.62 | 0
| 1 | 3 | 2.4873 | 3.3561 | 4.19 | 3.16 | 3.06 | 3.06 | 3.43 | 3.49 | 3.0136 | 3.79 | 2.93 | 2.7804 | 3.24 | 3.02 | 2.65 | 16
| 1 | 4 | 2.6236 | 3.4073 | 4.21 | 3.05 | 2.97 | 3.27 | 3.41 | 3.35 | 3.0652 | 3.75 | 2.96 | 2.7437 | 3.31 | 3.04 | 2.61 | 3.112
| 1 | 5 | 2.5986 | 3.3093 | 4.32 | 3.15 | 2.97 | 3.33 | 3.39 | 3.33 | 2.9213 | 3.8 | 2.99 | 2.7208 | 3.26 | 3.07 | 2.69 | 80
| 1 | 6 | 2.5211 | 3.3926 | 4.19 | 3.1 | 3.02 | 3.28 | 3.49 | 3.43 | 3.0424 | 3.79 | 2.92 | 2.5508 | 3.23 | 3.02 | 2.65
| Mean |  | 2.5452 | 3.3784 | 4.2 | 3.1283 | 3.0167 | 3.2067 | 3.425 | 3.4017 | 3.0675 | 3.78 | 2.9467 | 2.7334 | 3.255 | 3.03 | 2.6533
| Median |  | 2.5262 | 3.3961 | 4.2 | 3.145 | 3.02 | 3.225 | 3.415 | 3.39 | 3.0538 | 3.785 | 2.945 | 2.7508 | 3.25 | 3.03 | 2.65
| Std Dev. |  | 0.053657 | 0.038763 | 0.15047 | 0.04535 | 0.040332 | 0.10424 | 0.034496 | 0.067057 | 0.10155 | 0.017889 | 0.028048 | 0.099297 | 0.030166 | 0.026833 | 0.036148
| Rel.Std.Dev. |  | 2.108% | 1.147% | 3.583% | 1.45% | 1.337% | 3.251% | 1.007% | 1.971% | 3.311% | 0.4732% | 0.9518% | 3.633% | 0.9268% | 0.8856% | 1.362%
| PDM3 |  | -18.21% | 8.562% | 34.96% | 0.5252% | -3.063% | 3.042% | 10.06% | 9.308% | -1.431% | 21.47% | -5.312% | -12.17% | 4.595% | -2.635% | -14.74%
| Z-Score (Absolute) |  | 1.57 | 0.67 | 2.89 | 0 | 0.3 | 0.21 | 0.8 | 0.74 | 0.16 | 1.76 | 0.49 | 1.06 | 0.34 | 0.27 | 1.28
| Analytical results for Cd in OREAS 230b (Certified Value 0.27 ppm) | [16 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.275 | < 0.5 | 0.26 | 0.27 | 0.27 | 0.25 | 0.26 | 0.26 | 0.24231 | 0.27 | 0.28 | 0.3 | 0.3 | 0.24 |  | 1
| 1 | 2 | 0.27 | < 0.5 | 0.27 | 0.27 | 0.28 | 0.26 | 0.29 | 0.26 | 0.24548 | 0.26 | 0.26 | 0.3 | 0.3 | 0.25 |  | 24
| 1 | 3 | 0.263 | < 0.5 | 0.24 | 0.28 | 0.26 | 0.26 | 0.27 | 0.28 | 0.22962 | 0.27 | 0.27 | 0.3 | 0.2 | 0.27 |  | 16
| 1 | 4 | 0.26 | < 0.5 | 0.3 | 0.27 | 0.27 | 0.26 | 0.26 | 0.26 | 0.24618 | 0.25 | 0.27 | 0.3 | 0.3 | 0.27 |  | 0.26691
| 1 | 5 | 0.293 | < 0.5 | 0.27 | 0.25 | 0.26 | 0.29 | 0.26 | 0.25 | 0.24002 | 0.28 | 0.28 | 0.3 | 0.2 | 0.26 |  | 81
| 1 | 6 |  | < 0.5 | 0.27 | 0.27 | 0.27 | 0.23 | 0.26 | 0.25 | 0.23385 | 0.28 | 0.27 | 0.3 | 0.3 | 0.24
| Mean |  | 0.2722 |  | 0.26833 | 0.26833 | 0.26833 | 0.25833 | 0.26667 | 0.26 | 0.23958 | 0.26833 | 0.27167 | 0.3 | 0.26667 | 0.255
| Median |  | 0.27 |  | 0.27 | 0.27 | 0.27 | 0.26 | 0.26 | 0.26 | 0.24116 | 0.27 | 0.27 | 0.3 | 0.3 | 0.255
| Std Dev. |  | 0.013027 |  | 0.019408 | 0.0098319 | 0.0075277 | 0.019408 | 0.012111 | 0.010954 | 0.0066058 | 0.01169 | 0.0075277 | 0 | 0.05164 | 0.013784
| Rel.Std.Dev. |  | 4.786% |  | 7.233% | 3.664% | 2.805% | 7.513% | 4.541% | 4.213% | 2.757% | 4.357% | 2.771% | 0% | 19.36% | 5.406%
| PDM3 |  | 1.982% |  | 0.5328% | 0.5328% | 0.5328% | -3.214% | -0.09158% | -2.589% | -10.24% | 0.5328% | 1.782% | 12.4% | -0.09158% | -4.463%
| Z-Score (Absolute) |  | 0.71 | 2.66 | 0.13 | 0.13 | 0.13 | 1.39 | 0.13 | 1.14 | 4.25 | 0.13 | 0.63 | NA | NA | 1.9
Results from laboratories 22 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 230b (Certified Value 6.15 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 6.524 | 6.1017 | 7.23 | 6.14 | 6.06 | 6.05 | 6.82 | 6.18 | 6.89 | 5.24 | 5.9 | 5 | 5.11 |  |  | 1
| 1 | 2 | 6.325 | 5.9998 | 6.98 | 6.49 | 6.16 | 6.11 | 6.96 | 6.19 | 6.9 | 5.26 | 5.8 | 5 | 5.12 |  |  | 25
| 1 | 3 | 6.381 | 5.9691 | 6.77 | 6.34 | 5.99 | 5.89 | 7.07 | 6.27 | 7.04 | 5.12 | 5.4 | 5 | 5.23 |  |  | 16
| 1 | 4 | 6.416 | 5.922 | 7.11 | 6.47 | 5.96 | 6.17 | 6.94 | 6.19 | 6.98 | 5.14 | 5.9 | 6 | 5.14 |  |  | 6.1488
| 1 | 5 | 6.379 | 5.7538 | 7.32 | 6.35 | 5.97 | 6.34 | 7.01 | 6.07 | 6.82 | 5.22 | 5 | 5 | 5.17 |  |  | 82
| 1 | 6 | 6 | 5.9112 | 6.98 | 6.46 | 5.95 | 6.15 | 6.88 | 5.94 | 6.97 | 5.13 | 5.4 | 5 | 5.19
| Mean |  | 6.3375 | 5.9429 | 7.065 | 6.375 | 6.015 | 6.1183 | 6.9467 | 6.14 | 6.9333 | 5.185 | 5.5667 | 5.1667 | 5.16
| Median |  | 6.38 | 5.9455 | 7.045 | 6.405 | 5.98 | 6.13 | 6.95 | 6.185 | 6.935 | 5.18 | 5.6 | 5 | 5.155
| Std Dev. |  | 0.17811 | 0.1152 | 0.19786 | 0.13157 | 0.081179 | 0.14811 | 0.089368 | 0.11696 | 0.078401 | 0.061887 | 0.36148 | 0.40825 | 0.045607
| Rel.Std.Dev. |  | 2.81% | 1.938% | 2.801% | 2.064% | 1.35% | 2.421% | 1.286% | 1.905% | 1.131% | 1.194% | 6.494% | 7.902% | 0.8839%
| PDM3 |  | 3.069% | -3.348% | 14.9% | 3.679% | -2.176% | -0.4953% | 12.98% | -0.1429% | 12.76% | -15.67% | -9.467% | -15.97% | -16.08%
| Z-Score (Absolute) |  | 0.35 | 0.31 | 1.56 | 0.41 | 0.19 | 0.02 | 1.36 | 0.02 | 1.34 | 1.58 | 0.94 | NA | 1.62
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 230b (Certified Value 32.6 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 30.309 | 31.9979 | 34.1 | 32.7 | 32.1 | 30.4 | 32.3 | 31.6 | 35.8902 | 32.7 | 34.6 | 33.5 | 27.1 | 28.3 |  | 1
| 1 | 2 | 31.721 | 31.6469 | 31.3 | 32.8 | 33.1 | 30.4 | 33.6 | 30.2 | 34.4869 | 32.8 | 34.1 | 33 | 28 | 27.8 |  | 26
| 1 | 3 | 30.351 | 31.4205 | 31.4 | 32 | 31.9 | 30.5 | 33.4 | 32 | 35.5049 | 33 | 34.4 | 33.5 | 28 | 28 |  | 16
| 1 | 4 | 31.198 | 32.814 | 32 | 32.9 | 32.8 | 30.7 | 32.1 | 32.2 | 35.7356 | 33.1 | 34.1 | 32.2 | 28 | 28.3 |  | 32.5505
| 1 | 5 | 33.159 | 31.192 | 31.8 | 33.6 | 31.8 | 30.8 | 34.2 | 29.5 | 34.8165 | 33.4 | 33.6 | 33.7 | 27.8 | 28.3 |  | 83
| 1 | 6 | 32 | 31.4601 | 33.3 | 32.9 | 32.4 | 30.9 | 32.3 | 30.9 | 35.1057 | 32.7 | 33.9 | 32.9 | 27.7 | 28.4
| Mean |  | 31.4563 | 31.7552 | 32.3167 | 32.8167 | 32.35 | 30.6167 | 32.9833 | 31.0667 | 35.2566 | 32.95 | 34.1167 | 33.1333 | 27.7667 | 28.1833
| Median |  | 31.4595 | 31.5535 | 31.9 | 32.85 | 32.25 | 30.6 | 32.85 | 31.25 | 35.3053 | 32.9 | 34.1 | 33.25 | 27.9 | 28.3
| Std Dev. |  | 1.0835 | 0.58446 | 1.1303 | 0.51153 | 0.51672 | 0.2137 | 0.86583 | 1.0652 | 0.54799 | 0.27386 | 0.35449 | 0.55377 | 0.35024 | 0.23166
| Rel.Std.Dev. |  | 3.444% | 1.841% | 3.498% | 1.559% | 1.597% | 0.698% | 2.625% | 3.429% | 1.554% | 0.8311% | 1.039% | 1.671% | 1.261% | 0.822%
| PDM3 |  | -3.362% | -2.443% | -0.7185% | 0.8176% | -0.6161% | -5.941% | 1.33% | -4.559% | 8.313% | 1.227% | 4.811% | 1.79% | -14.7% | -13.42%
| Z-Score (Absolute) |  | 0.71 | 0.46 | 0.01 | 0.39 | 0.01 | 1.38 | 0.52 | 1.02 | 2.35 | 0.5 | 1.43 | 0.64 | 3.67 | 3.34
| Analytical results for Cr in OREAS 230b (Certified Value 21.6 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 19.732 | 22.9281 | 24 | 20 | 20 | 20 | 22 | 22 | 25.7131 | 23 | 22 | 22 | 21.1 |  |  | 1
| 1 | 2 | 21.854 | 22.5766 | 23 | 20 | 21 | 20 | 21 | 22 | 24.9114 | 23 | 22 | 22 | 20.7 |  |  | 27
| 1 | 3 | 20.256 | 21.9493 | 23 | 20 | 21 | 21 | 22 | 22 | 25.8498 | 23 | 21 | 23 | 21.1 |  |  | 16
| 1 | 4 | 21.313 | 22.7129 | 24 | 19 | 20 | 21 | 22 | 21 | 25.9008 | 24 | 21 | 22 | 20.8 |  |  | 21.5785
| 1 | 5 | 21.689 | 21.9904 | 24 | 20 | 20 | 21 | 21 | 21 | 25.287 | 24 | 21 | 21 | 21.4 |  |  | 84
| 1 | 6 | 21 | 22.353 | 24 | 20 | 20 | 21 | 22 | 22 | 25.4036 | 24 | 20 | 22 | 21.2
| Mean |  | 20.974 | 22.4184 | 23.6667 | 19.8333 | 20.3333 | 20.6667 | 21.6667 | 21.6667 | 25.5109 | 23.5 | 21.1667 | 22 | 21.05
| Median |  | 21.1565 | 22.4648 | 24 | 20 | 20 | 21 | 22 | 22 | 25.5584 | 23.5 | 21 | 22 | 21.1
| Std Dev. |  | 0.83178 | 0.39471 | 0.5164 | 0.40825 | 0.5164 | 0.5164 | 0.5164 | 0.5164 | 0.3818 | 0.54772 | 0.75277 | 0.63246 | 0.25884
| Rel.Std.Dev. |  | 3.966% | 1.761% | 2.182% | 2.058% | 2.54% | 2.499% | 2.383% | 2.383% | 1.497% | 2.331% | 3.556% | 2.875% | 1.23%
| PDM3 |  | -2.802% | 3.892% | 9.677% | -8.088% | -5.771% | -4.226% | 0.4084% | 0.4084% | 18.22% | 8.905% | -1.909% | 1.953% | -2.449%
| Z-Score (Absolute) |  | 0.62 | 0.67 | 1.79 | 1.64 | 1.2 | 0.9 | 0 | 0 | 3.45 | 1.64 | 0.45 | 0.3 | 0.55
| Analytical results for Cs in OREAS 230b (Certified Value 0.34 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.364 | 0.31895 | 0.32 | 0.33 | 0.37 | 0.35 | 0.34 | 0.33 | 0.34 | 0.36 | < 1 | 0.3 | 0.32 |  |  | 1
| 1 | 2 | 0.364 | 0.33634 | 0.32 | 0.33 | 0.39 | 0.33 | 0.35 | 0.32 | 0.34 | 0.35 | < 1 | 0.3 | 0.32 |  |  | 28
| 1 | 3 | 0.434 | 0.31909 | 0.3 | 0.34 | 0.37 | 0.34 | 0.35 | 0.34 | 0.34 | 0.35 | < 1 | 0.3 | 0.33 |  |  | 16
| 1 | 4 | 0.351 | 0.32387 | 0.29 | 0.33 | 0.37 | 0.34 | 0.34 | 0.33 | 0.34 | 0.35 | < 1 | 0.3 | 0.33 |  |  | 0.33828
| 1 | 5 | 0.321 | 0.33858 | 0.33 | 0.34 | 0.36 | 0.33 | 0.34 | 0.33 | 0.34 | 0.35 | < 1 | 0.3 | 0.33 |  |  | 85
| 1 | 6 |  | 0.32988 | 0.31 | 0.34 | 0.37 | 0.35 | 0.34 | 0.32 | 0.34 | 0.33 | < 1 | 0.3 | 0.32
| Mean |  | 0.3668 | 0.32778 | 0.31167 | 0.335 | 0.37167 | 0.34 | 0.34333 | 0.32833 | 0.34 | 0.34833 |  | 0.3 | 0.325
| Median |  | 0.364 | 0.32687 | 0.315 | 0.335 | 0.37 | 0.34 | 0.34 | 0.33 | 0.34 | 0.35 |  | 0.3 | 0.325
| Std Dev. |  | 0.041469 | 0.0085241 | 0.01472 | 0.0054772 | 0.0098319 | 0.0089443 | 0.005164 | 0.0075277 | 0 | 0.0098319 |  | 0 | 0.0054772
| Rel.Std.Dev. |  | 11.31% | 2.601% | 4.723% | 1.635% | 2.645% | 2.631% | 1.504% | 2.293% | 0% | 2.823% |  | 0% | 1.685%
| PDM3 |  | 8.43% | -3.104% | -7.868% | -0.9706% | 9.868% | 0.5075% | 1.493% | -2.941% | 0.5075% | 2.971% |  | -11.32% | -3.927%
| Z-Score (Absolute) |  | 1.51 | 0.69 | 1.6 | 0.28 | 1.79 | 0 | 0.19 | 0.66 | 0 | 0.47 | 9.03 | NA | 0.85
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 230b (Certified Value 179 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 0
| 1 | 1 | 170.148 | 181.3046 | 182 | 179 | 180 | 184 | 182 | 177.5 | 216.6798 | 172 | 178.4 | 185 | 191.5 | 199.6 | 179 | 1
| 1 | 2 | 179.533 | 182.599 | 166 | 182 | 183 | 184 | 181.5 | 175 | 213.4747 | 172 | 174.5 | 175 | 194.1 | 199.7 | 175 | 29
| 1 | 3 | 175.915 | 179.6047 | 168 | 180.5 | 180 | 178 | 180.5 | 179.5 | 203.9022 | 173 | 172.2 | 180 | 192.3 | 199.2 | 177 | 16
| 1 | 4 | 178.599 | 183.1786 | 168 | 180 | 179 | 186.5 | 180 | 181 | 207.9726 | 172 | 172.2 | 174 | 187.9 | 196 | 175 | 178.702
| 1 | 5 | 182.205 | 178.5673 | 176 | 179 | 176 | 187.5 | 182 | 171 | 202.5807 | 171 | 171.9 | 183 | 184.2 | 194.7 | 178 | 86
| 1 | 6 | 180 | 181.1034 | 173 | 182 | 178.5 | 186.5 | 186 | 176.5 | 206.4127 | 174 | 169.8 | 169 | 185 | 196.2 | 177
| Mean |  | 177.7333 | 181.0596 | 172.1667 | 180.4167 | 179.4167 | 184.4167 | 182 | 176.75 | 208.5038 | 172.3333 | 173.1667 | 177.6667 | 189.1667 | 197.5667 | 176.8333
| Median |  | 179.066 | 181.204 | 170.5 | 180.25 | 179.5 | 185.25 | 181.75 | 177 | 207.1926 | 172 | 172.2 | 177.5 | 189.7 | 197.7 | 177
| Std Dev. |  | 4.2418 | 1.7464 | 6.08 | 1.3571 | 2.2895 | 3.4557 | 2.1213 | 3.532 | 5.5221 | 1.0328 | 2.9656 | 6.0553 | 4.08 | 2.186 | 1.6021
| Rel.Std.Dev. |  | 2.387% | 0.9646% | 3.531% | 0.7522% | 1.276% | 1.874% | 1.166% | 1.998% | 2.648% | 0.5993% | 1.713% | 3.408% | 2.157% | 1.106% | 0.906%
| PDM3 |  | -0.5421% | 1.319% | -3.657% | 0.9595% | 0.3999% | 3.198% | 1.846% | -1.092% | 16.68% | -3.564% | -3.098% | -0.5794% | 5.856% | 10.56% | -1.046%
| Z-Score (Absolute) |  | 0.43 | 0.42 | 1.83 | 0.25 | 0 | 1.26 | 0.65 | 0.67 | 7.36 | 1.79 | 1.58 | 0.44 | 2.47 | 4.59 | 0.65
| Analytical results for Dy in OREAS 230b (Indicative Value 2.46 ppm) | [16 empty] | Indicative
Dy | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 2.4359; 1.06: 2.6 | 1
| 1 | 2 | 1.05: 2.434; 1.06: 2.5 | 4
| 1 | 3 | 1.05: 2.3289; 1.06: 2.5 | 16
| 1 | 4 | 1.05: 2.3471; 1.06: 2.4 | 2.4568
| 1 | 5 | 1.05: 2.3159; 1.06: 2.5 | 10
| 1 | 6 | 1.05: 2.4196; 1.06: 2.7
| Mean |  | 1.05: 2.3802; 1.06: 2.5333
| Median |  | 1.05: 2.3833; 1.06: 2.5
| Std Dev. |  | 1.05: 0.055535; 1.06: 0.10328
| Rel.Std.Dev. |  | 1.05: 2.333%; 1.06: 4.077%
| PDM3 |  | 1.05: -3.116%; 1.06: 3.116%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Er in OREAS 230b (Indicative Value 1.56 ppm) | [16 empty] | Indicative
Er | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 1.512; 1.06: 1.6 | 1
| 1 | 2 | 1.05: 1.446; 1.06: 1.6 | 5
| 1 | 3 | 1.05: 1.4796; 1.06: 1.6 | 16
| 1 | 4 | 1.05: 1.4715; 1.06: 1.6 | 1.5558
| 1 | 5 | 1.05: 1.4768; 1.06: 1.7 | 11
| 1 | 6 | 1.05: 1.4836; 1.06: 1.7
| Mean |  | 1.05: 1.4783; 1.06: 1.6333
| Median |  | 1.05: 1.4782; 1.06: 1.6
| Std Dev. |  | 1.05: 0.021264; 1.06: 0.05164
| Rel.Std.Dev. |  | 1.05: 1.438%; 1.06: 3.162%
| PDM3 |  | 1.05: -4.984%; 1.06: 4.984%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Eu in OREAS 230b (Indicative Value 0.53 ppm) | [16 empty] | Indicative
Eu | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 0.48713; 1.06: 0.6 | 1
| 1 | 2 | 1.05: 0.4964; 1.06: 0.5 | 6
| 1 | 3 | 1.05: 0.47296; 1.06: 0.5 | 16
| 1 | 4 | 1.05: 0.47535; 1.06: 0.6 | 0.52586
| 1 | 5 | 1.05: 0.48785; 1.06: 0.6 | 12
| 1 | 6 | 1.05: 0.49066; 1.06: 0.6
| Mean |  | 1.05: 0.48506; 1.06: 0.56667
| Median |  | 1.05: 0.48749; 1.06: 0.6
| Std Dev. |  | 1.05: 0.0090844; 1.06: 0.05164
| Rel.Std.Dev. |  | 1.05: 1.873%; 1.06: 9.113%
| PDM3 |  | 1.05: -7.759%; 1.06: 7.759%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Fe in OREAS 230b (Certified Value 6.38 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 5.7049 | 6.6581 | 6.57 | 6.34 | 6.33 | 6.11 | 6.56 | 6.14 | 6.884 | 6.65 | 6.3 | 6.5375 | 6.29 | 6.2 | 6.66 | 1
| 1 | 2 | 6.2109 | 6.6429 | 6.35 | 6.28 | 6.39 | 6.18 | 6.62 | 6.04 | 6.6253 | 6.71 | 6.28 | 6.2349 | 6.29 | 6.33 | 6.5 | 0
| 1 | 3 | 5.9638 | 6.6017 | 6.19 | 6.28 | 6.33 | 6.44 | 6.68 | 6.25 | 6.8236 | 6.74 | 6.28 | 6.4092 | 6.32 | 6.24 | 6.62 | 16
| 1 | 4 | 6.2401 | 6.6568 | 6.37 | 6.25 | 6.24 | 6.24 | 6.59 | 6.34 | 6.8762 | 6.69 | 6.32 | 6.2583 | 6.28 | 6.33 | 6.39 | 6.3768
| 1 | 5 | 6.3573 | 6.5661 | 6.53 | 6.28 | 6.24 | 6.36 | 6.52 | 6.05 | 6.6851 | 6.73 | 6.36 | 6.2877 | 6.29 | 6.36 | 6.48 | 87
| 1 | 6 | 6.171 | 6.5764 | 6.77 | 6.29 | 6.28 | 6.29 | 6.72 | 6.03 | 6.7413 | 6.78 | 6.29 | 5.9501 | 6.23 | 6.26 | 6.47
| Mean |  | 6.108 | 6.617 | 6.4633 | 6.2867 | 6.3017 | 6.27 | 6.615 | 6.1417 | 6.7726 | 6.7167 | 6.305 | 6.2796 | 6.2833 | 6.2867 | 6.52
| Median |  | 6.191 | 6.6223 | 6.45 | 6.28 | 6.305 | 6.265 | 6.605 | 6.095 | 6.7825 | 6.72 | 6.295 | 6.273 | 6.29 | 6.295 | 6.49
| Std Dev. |  | 0.23556 | 0.041029 | 0.20304 | 0.029439 | 0.059133 | 0.12 | 0.074766 | 0.12828 | 0.10591 | 0.044572 | 0.030822 | 0.19718 | 0.029439 | 0.062503 | 0.101
| Rel.Std.Dev. |  | 3.857% | 0.62% | 3.141% | 0.4683% | 0.9384% | 1.914% | 1.13% | 2.089% | 1.564% | 0.6636% | 0.4889% | 3.14% | 0.4685% | 0.9942% | 1.549%
| PDM3 |  | -4.215% | 3.767% | 1.357% | -1.414% | -1.178% | -1.675% | 3.735% | -3.687% | 6.206% | 5.33% | -1.126% | -1.524% | -1.466% | -1.414% | 2.246%
| Z-Score (Absolute) |  | 0.82 | 1.33 | 0.68 | 0.06 | 0 | 0.13 | 1.32 | 0.67 | 1.98 | 1.75 | 0.01 | 0.09 | 0.08 | 0.06 | 0.92
| Analytical results for Ga in OREAS 230b (Certified Value 11.1 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 10.2483 | 11.9 | 10.95 | 11.2 | 10.95 | 12.25 | 11.1 | 14.7682 | 12.7 | 10.7 | 14.4 | 11 | 9.12 |  |  | 1
| 1 | 2 | 10.2762 | 11.4 | 11.05 | 12.25 | 11.35 | 11.7 | 10.55 | 13.3867 | 12.7 | 10.8 | 14.7 | 11 | 8.88 |  |  | 33
| 1 | 3 | 10.5094 | 11.4 | 11.1 | 11.35 | 10.65 | 12.2 | 11.6 | 15.1538 | 12.8 | 10.6 | 14.9 | 11 | 8.96 |  |  | 16
| 1 | 4 | 10.4361 | 11.4 | 10.8 | 11.35 | 11.2 | 12 | 11.5 | 15.0598 | 12.6 | 10.6 | 14.8 | 11 | 9.07 |  |  | 11.0779
| 1 | 5 | 10.7006 | 11.6 | 11.1 | 11.3 | 11.35 | 12.05 | 9.3 | 15.2718 | 13.1 | 10.5 | 14.7 | 11 | 9.12 |  |  | 88
| 1 | 6 | 10.4239 | 11.6 | 11 | 11.6 | 11.15 | 11.9 | 10.6 | 15.2506 | 12.6 | 10.4 | 14.4 | 11 | 9.08
| Mean |  | 10.4324 | 11.55 | 11 | 11.5083 | 11.1083 | 12.0167 | 10.775 | 14.8151 | 12.75 | 10.6 | 14.65 | 11 | 9.0383
| Median |  | 10.43 | 11.5 | 11.025 | 11.35 | 11.175 | 12.025 | 10.85 | 15.1068 | 12.7 | 10.6 | 14.7 | 11 | 9.075
| Std Dev. |  | 0.16507 | 0.19748 | 0.11402 | 0.38654 | 0.2691 | 0.20166 | 0.84484 | 0.72325 | 0.18708 | 0.14142 | 0.20736 | 0 | 0.097245
| Rel.Std.Dev. |  | 1.582% | 1.71% | 1.037% | 3.359% | 2.423% | 1.678% | 7.841% | 4.882% | 1.467% | 1.334% | 1.415% | 0% | 1.076%
| PDM3 |  | -5.827% | 4.262% | -0.7033% | 3.885% | 0.2747% | 8.474% | -2.734% | 33.74% | 15.09% | -4.314% | 32.25% | -0.7033% | -18.41%
| Z-Score (Absolute) |  | 0.83 | 0.23 | 0.29 | 0.19 | 0.19 | 0.67 | 0.51 | 3.34 | 1.37 | 0.67 | 3.18 | NA | 2.16
| Analytical results for Gd in OREAS 230b (Indicative Value 1.92 ppm) | [16 empty] | Indicative
Gd | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 1.835; 1.06: 2 | 1
| 1 | 2 | 1.05: 1.8676; 1.06: 2 | 7
| 1 | 3 | 1.05: 1.822; 1.06: 1.9 | 16
| 1 | 4 | 1.05: 1.8444; 1.06: 2 | 1.9206
| 1 | 5 | 1.05: 1.7593; 1.06: 2.1 | 13
| 1 | 6 | 1.05: 1.8189; 1.06: 2.1
| Mean |  | 1.05: 1.8245; 1.06: 2.0167
| Median |  | 1.05: 1.8285; 1.06: 2
| Std Dev. |  | 1.05: 0.036458; 1.06: 0.075277
| Rel.Std.Dev. |  | 1.05: 1.998%; 1.06: 3.733%
| PDM3 |  | 1.05: -5.002%; 1.06: 5.002%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Ge in OREAS 230b (Certified Value 0.17 ppm) | [16 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.22 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  |  | 2
| 1 | 1 | 0.29292 | 0.1 | 0.15 | 0.15 | 0.16 | 0.15 | 0.22 | 7.0611 | 0.2 | < 0.5 | 1.7 | 0.12 |  |  |  | 1
| 1 | 2 | 0.28293 | 0.1 | 0.15 | 0.15 | 0.16 | 0.15 | 0.21 | 5.4996 | 0.2 | < 0.5 | 1.7 | 0.11 |  |  |  | 1
| 1 | 3 | 0.24147 | 0.1 | 0.15 | 0.17 | 0.2 | 0.15 | 0.22 | 6.4739 | 0.2 | < 0.5 | 1.7 | 0.12 |  |  |  | 16
| 1 | 4 | 0.26456 | 0.1 | 0.14 | 0.17 | 0.18 | 0.16 | 0.2 | 6.4665 | 0.2 | < 0.5 | 1.6 | 0.12 |  |  |  | 0.17349
| 1 | 5 | 0.22512 | 0.1 | 0.15 | 0.16 | 0.18 | 0.16 | 0.2 | 6.4488 | 0.2 | < 0.5 | 1.6 | 0.11 |  |  |  | 89
| 1 | 6 | 0.19946 | 0.1 | 0.16 | 0.17 | 0.15 | 0.14 | 0.22 | 6.221 | 0.2 | < 0.5 | 1.6 | 0.12
| Mean |  | 0.25108 | 0.1 | 0.15 | 0.16167 | 0.17167 | 0.15167 | 0.21167 | 6.3618 | 0.2 |  | 1.65 | 0.11667
| Median |  | 0.25301 | 0.1 | 0.15 | 0.165 | 0.17 | 0.15 | 0.215 | 6.4576 | 0.2 |  | 1.65 | 0.12
| Std Dev. |  | 0.035714 | 1.5202e-17 | 0.0063246 | 0.0098319 | 0.018348 | 0.0075277 | 0.0098319 | 0.5066 | 3.0405e-17 |  | 0.054772 | 0.005164
| Rel.Std.Dev. |  | 14.22% | 1.52e-14% | 4.216% | 6.082% | 10.69% | 4.963% | 4.645% | 7.963% | 1.52e-14% |  | 3.32% | 4.426%
| PDM3 |  | 44.72% | -42.36% | -13.54% | -6.813% | -1.049% | -12.58% | 22.01% | 3567% | 15.28% |  | 851.1% | -32.75%
| Z-Score (Absolute) |  | 1.34 | NA | 0.37 | 0.17 | 0 | 0.34 | 0.67 | 104.35 | NA | 1.32 | NA | 0.93
| Analytical results for Hf in OREAS 230b (Certified Value 0.51 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | 0.55019 | 0.4 | 0.51 | 0.51 | 0.59 | 0.52 | 0.5 | 0.048983 | 0.82 | 0.46 | 0.4 | 0.43 |  |  |  | 1
| 1 | 2 | 0.54368 | 0.3 | 0.53 | 0.53 | 0.49 | 0.53 | 0.5 | 0.075025 | 0.8 | 0.47 | 0.4 | 0.44 |  |  |  | 19
| 1 | 3 | 0.54872 | 0.3 | 0.55 | 0.54 | 0.49 | 0.53 | 0.5 | 0.049961 | 0.81 | 0.45 | 0.5 | 0.44 |  |  |  | 16
| 1 | 4 | 0.55401 | 0.4 | 0.52 | 0.5 | 0.56 | 0.54 | 0.5 | 0.045987 | 0.81 | 0.46 | 0.4 | 0.43 |  |  |  | 0.50893
| 1 | 5 | 0.546 | 0.4 | 0.53 | 0.51 | 0.62 | 0.52 | 0.51 | 0.066163 | 0.81 | 0.44 | 0.5 | 0.44 |  |  |  | 90
| 1 | 6 | 0.55593 | 0.3 | 0.55 | 0.52 | 0.55 | 0.54 | 0.51 | 0.062439 | 0.8 | 0.44 | 0.4 | 0.43
| Mean |  | 0.54975 | 0.35 | 0.53167 | 0.51833 | 0.55 | 0.53 | 0.50333 | 0.058093 | 0.80833 | 0.45333 | 0.43333 | 0.435
| Median |  | 0.54945 | 0.35 | 0.53 | 0.515 | 0.555 | 0.53 | 0.5 | 0.0562 | 0.81 | 0.455 | 0.4 | 0.435
| Std Dev. |  | 0.0046554 | 0.054772 | 0.016021 | 0.01472 | 0.052536 | 0.0089443 | 0.005164 | 0.011544 | 0.0075277 | 0.012111 | 0.05164 | 0.0054772
| Rel.Std.Dev. |  | 0.8468% | 15.65% | 3.013% | 2.84% | 9.552% | 1.688% | 1.026% | 19.87% | 0.9313% | 2.671% | 11.92% | 1.259%
| PDM3 |  | 8.022% | -31.23% | 4.468% | 1.848% | 8.07% | 4.141% | -1.099% | -88.59% | 58.83% | -10.92% | -14.85% | -14.53%
| Z-Score (Absolute) |  | 0.67 | NA | 0.2 | 0.15 | 0.68 | 0.15 | 0.55 | 12.22 | 7.45 | 1.86 | NA | 2.34
Results from laboratories 7 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 230b (Indicative Value 0.031 ppm) | [16 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.19 | 1.21 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.181 | < 1 | 0.06 | 0.02 | 0.01 | 0.02 | 0.02 | 0.03 | 0.072046 | 0.02 | 0.14 | 0.34 | 0.15 |  |  | 1
| 1 | 2 | 0.222 | < 1 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.054282 | 0.03 | 0.12 | 0.342 | 0.15 |  |  | 8
| 1 | 3 | 0.186 | < 1 | 0.05 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.06959 | 0.02 | 0.12 | 0.363 | 0.15 |  |  | 16
| 1 | 4 | 0.134 | < 1 | 0.06 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.076024 | 0.03 | 0.11 | 0.344 | 0.15 |  |  | 0.030545
| 1 | 5 | 0.184 | < 1 | 0.05 | 0.02 | 0.02 | 0.01 | 0.03 | 0.02 | 0.064728 | 0.02 | 0.15 | 0.357 | 0.15 |  |  | 14
| 1 | 6 |  | < 1 | 0.05 | 0.02 | 0.01 | 0.04 | 0.02 | 0.02 | 0.059503 | 0.02 | 0.11 | 0.395 | 0.15
| Mean |  | 0.1814 |  | 0.05 | 0.02 | 0.016667 | 0.025 | 0.021667 | 0.021667 | 0.066029 | 0.023333 | 0.125 | 0.35683 | 0.15
| Median |  | 0.184 |  | 0.05 | 0.02 | 0.02 | 0.025 | 0.02 | 0.02 | 0.067159 | 0.02 | 0.12 | 0.3505 | 0.15
| Std Dev. |  | 0.031318 |  | 0.010954 | 0 | 0.005164 | 0.010488 | 0.0040825 | 0.0040825 | 0.0081372 | 0.005164 | 0.016432 | 0.020798 | 0
| Rel.Std.Dev. |  | 17.26% |  | 21.91% | 0% | 30.98% | 41.95% | 18.84% | 18.84% | 12.32% | 22.13% | 13.15% | 5.829% | 0%
| PDM3 |  | 493.9% |  | 63.69% | -34.52% | -45.44% | -18.15% | -29.07% | -29.07% | 116.2% | -23.61% | 309.2% | 1068% | 391.1%
| Z-Score (Absolute) |  | 2.95 | 10.11 | 0 | 0.67 | 0.75 | 0.56 | 0.64 | 0.64 | 0.36 | 0.6 | 1.69 | 6.9 | 2.25
| Analytical results for Ho in OREAS 230b (Indicative Value 0.51 ppm) | [16 empty] | Indicative
Ho | Round | Replicate | 1.06: Lab; 1.07: Lab |  |  |  | 1
| No | No | 1.06: 1.06; 1.07: 1.07 |  |  |  | ppm
|  |  | 1.06: AR*MS; 1.07: AR*OES/MS |  |  |  | 2
|  |  | 1.06: 0.2g; 1.07: 0.5g |  |  |  | 2
| 1 | 1 | 1.06: 0.50285; 1.07: 0.6 |  |  |  | 1
| 1 | 2 | 1.06: 0.50503; 1.07: 0.5 |  |  |  | 3
| 1 | 3 | 1.06: 0.50246; 1.07: 0.5 |  |  |  | 16
| 1 | 4 | 1.06: 0.50133; 1.07: 0.5 |  |  |  | 0.51494
| 1 | 5 | 1.06: 0.47787; 1.07: 0.6 |  |  |  | 9
| 1 | 6 | 1.06: 0.48977; 1.07: 0.5
| Mean |  | 1.06: 0.49655; 1.07: 0.53333
| Median |  | 1.06: 0.5019; 1.07: 0.5
| Std Dev. |  | 1.06: 0.010622; 1.07: 0.05164
| Rel.Std.Dev. |  | 1.06: 2.139%; 1.07: 9.682%
| PDM3 |  | 1.06: -3.571%; 1.07: 3.571%
| Z-Score (Absolute) |  | 1.06: 0.67; 1.07: 0.67
| Analytical results for In in OREAS 230b (Certified Value 0.03 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | < 0.1 | 0.03 | 0.032 | 0.029 | 0.029 | 0.033 | 0.033 | 0.028183 | < 0.05 | 0.03 | < 0.5 | 0.03 | 0.024 |  |  | 1
| 1 | 2 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033 | 0.032 | 0.026783 | < 0.05 | 0.03 | < 0.5 | 0.03 | 0.024 |  |  | 21
| 1 | 3 | < 0.1 | 0.03 | 0.032 | 0.027 | 0.031 | 0.032 | 0.033 | 0.028627 | < 0.05 | 0.03 | < 0.5 | 0.03 | 0.023 |  |  | 16
| 1 | 4 | < 0.1 | 0.04 | 0.031 | 0.028 | 0.03 | 0.033 | 0.03 | 0.041604 | < 0.05 | 0.03 | < 0.5 | 0.03 | 0.024 |  |  | 0.030191
| 1 | 5 | < 0.1 | 0.05 | 0.031 | 0.029 | 0.028 | 0.033 | 0.031 | 0.028012 | < 0.05 | 0.03 | < 0.5 | 0.03 | 0.025 |  |  | 91
| 1 | 6 | < 0.1 | 0.03 | 0.03 | 0.027 | 0.03 | 0.033 | 0.029 | 0.02948 | < 0.05 | 0.02 | < 0.5 | 0.03 | 0.024
| Mean |  |  | 0.035 | 0.031 | 0.028333 | 0.029667 | 0.032833 | 0.031333 | 0.030448 |  | 0.028333 |  | 0.03 | 0.024
| Median |  |  | 0.03 | 0.031 | 0.0285 | 0.03 | 0.033 | 0.0315 | 0.028405 |  | 0.03 |  | 0.03 | 0.024
| Std Dev. |  |  | 0.0083666 | 0.00089443 | 0.0012111 | 0.0010328 | 0.00040825 | 0.001633 | 0.0055354 |  | 0.0040825 |  | 0 | 0.00063246
| Rel.Std.Dev. |  |  | 23.9% | 2.885% | 4.274% | 3.481% | 1.243% | 5.212% | 18.18% |  | 14.41% |  | 0% | 2.635%
| PDM3 |  |  | 15.93% | 2.68% | -6.152% | -1.736% | 8.753% | 3.784% | 0.8526% |  | -6.152% |  | -0.6319% | -20.51%
| Z-Score (Absolute) |  | 6.23 | 1.45 | 0.18 | 0.67 | 0.25 | 0.76 | 0.28 | 0 | 1.74 | 0.67 | 70 | 0.14 | 2.06
| Analytical results for K in OREAS 230b (Certified Value 0.05 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.048585 | 0.050734 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.082594 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 |  | 1
| 1 | 2 | 0.051149 | 0.050754 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.081979 | 0.051 | 0.05 | 0.05 | 0.05 | 0.04 |  | 0
| 1 | 3 | 0.049686 | 0.052275 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.077536 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 |  | 16
| 1 | 4 | 0.049369 | 0.051239 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.078962 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 |  | 0.050052
| 1 | 5 | 0.051792 | 0.049348 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.076586 | 0.051 | 0.06 | 0.05 | 0.05 | 0.05 |  | 92
| 1 | 6 | 0.0495 | 0.050585 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.078278 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05
| Mean |  | 0.050014 | 0.050822 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.079323 | 0.050333 | 0.053333 | 0.05 | 0.05 | 0.048333
| Median |  | 0.049593 | 0.050744 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.07862 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
| Std Dev. |  | 0.0012064 | 0.00095127 | 0 | 7.6012e-18 | 7.6012e-18 | 7.6012e-18 | 7.6012e-18 | 7.6012e-18 | 0.0024354 | 0.0005164 | 0.005164 | 7.6012e-18 | 7.6012e-18 | 0.0040825
| Rel.Std.Dev. |  | 2.412% | 1.872% | 0% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 3.07% | 1.026% | 9.682% | 1.52e-14% | 1.52e-14% | 8.447%
| PDM3 |  | -0.07749% | 1.538% | -20.08% | -0.1045% | -0.1045% | -0.1045% | -0.1045% | -0.1045% | 58.48% | 0.5615% | 6.555% | -0.1045% | -0.1045% | -3.434%
| Z-Score (Absolute) |  | 1.35 | 82.05 | 997.74 | 0 | 0 | 0 | 0 | 0 | 2925.65 | 33.26 | 332.58 | 0 | 0 | 166.29
| Analytical results for La in OREAS 230b (Certified Value 2.34 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | 2.339 | 2.9 | 2.4 | 2.4 | 2.5 | 2.5 | 2.5 | 2.93 | 2.2 | 4.8 | 2 | 2 |  |  |  | 1
| 1 | 2 | 2.367 | 2.8 | 2.4 | 2.4 | 2.4 | 2.5 | 2.4 | 2.9 | 2.2 | 4.8 | 2 | 2 |  |  |  | 22
| 1 | 3 | 2.303 | 2.8 | 2.4 | 2.4 | 2.4 | 2.6 | 2.5 | 2.92 | 2.1 | 4.8 | 2 | 2.1 |  |  |  | 16
| 1 | 4 | 2.3966 | 2.9 | 2.4 | 2.4 | 2.5 | 2.5 | 2.4 | 2.83 | 2.1 | 4.8 | 2 | 2 |  |  |  | 2.3379
| 1 | 5 | 2.2632 | 2.9 | 2.4 | 2.4 | 2.5 | 2.5 | 2.4 | 2.91 | 2.1 | 4.7 | 2 | 2 |  |  |  | 93
| 1 | 6 | 2.3504 | 2.9 | 2.4 | 2.4 | 2.5 | 2.5 | 2.4 | 2.86 | 2.1 | 4.7 | 2 | 2
| Mean |  | 2.3365 | 2.8667 | 2.4 | 2.4 | 2.4667 | 2.5167 | 2.4333 | 2.8917 | 2.1333 | 4.7667 | 2 | 2.0167
| Median |  | 2.3447 | 2.9 | 2.4 | 2.4 | 2.5 | 2.5 | 2.4 | 2.905 | 2.1 | 4.8 | 2 | 2
| Std Dev. |  | 0.047427 | 0.05164 | 0 | 0 | 0.05164 | 0.040825 | 0.05164 | 0.038687 | 0.05164 | 0.05164 | 0 | 0.040825
| Rel.Std.Dev. |  | 2.03% | 1.801% | 0% | 0% | 2.094% | 1.622% | 2.122% | 1.338% | 2.421% | 1.083% | 0% | 2.024%
| PDM3 |  | -0.05848% | 22.62% | 2.656% | 2.656% | 5.508% | 7.646% | 4.082% | 23.69% | -8.75% | 103.9% | -14.45% | -13.74%
| Z-Score (Absolute) |  | 0.67 | 3.02 | 0.23 | 0.23 | 0.23 | 0.58 | 0 | 3.19 | 2.09 | 16.25 | NA | 2.9
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 230b (Certified Value 9.22 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 9.043 | 9.5783 | 8.1 | 9.2 | 9.5 | 9.2 | 9.2 | 9.2 | 12.7446 | 8.6 | 9 | 10.3 | 8.4 | 7.4 |  | 1
| 1 | 2 | 8.812 | 9.5406 | 7.8 | 9.1 | 9.6 | 10.7 | 9.2 | 9.1 | 12.0671 | 8.8 | 9 | 10.7 | 9.6 | 7.4 |  | 0
| 1 | 3 | 9.198 | 9.6666 | 8 | 9.3 | 9.5 | 9.4 | 9.2 | 9.4 | 11.811 | 8.7 | 10 | 9.9 | 9.2 | 7.7 |  | 16
| 1 | 4 | 9.023 | 9.4082 | 8.1 | 9.1 | 9.4 | 9.5 | 9.2 | 9.5 | 11.9248 | 8.7 | 9 | 10.3 | 9.2 | 7.8 |  | 9.2157
| 1 | 5 | 9.26 | 9.3445 | 8.1 | 9.2 | 9.3 | 9.7 | 9.1 | 9.1 | 12.538 | 8.6 | 9 | 10.4 | 9.5 | 7.9 |  | 94
| 1 | 6 | 9 | 9.3918 | 8.1 | 9.2 | 9.4 | 9.6 | 9.4 | 9.1 | 13.4721 | 8.8 | 9 | 10.5 | 8.9 | 8
| Mean |  | 9.056 | 9.4883 | 8.0333 | 9.1833 | 9.45 | 9.6833 | 9.2167 | 9.2333 | 12.4263 | 8.7 | 9.1667 | 10.35 | 9.1333 | 7.7
| Median |  | 9.033 | 9.4744 | 8.1 | 9.2 | 9.45 | 9.55 | 9.2 | 9.15 | 12.3026 | 8.7 | 9 | 10.35 | 9.2 | 7.75
| Std Dev. |  | 0.15856 | 0.12573 | 0.12111 | 0.075277 | 0.10488 | 0.52694 | 0.098319 | 0.17512 | 0.62716 | 0.089443 | 0.40825 | 0.26646 | 0.43665 | 0.25298
| Rel.Std.Dev. |  | 1.751% | 1.325% | 1.508% | 0.8197% | 1.11% | 5.442% | 1.067% | 1.897% | 5.047% | 1.028% | 4.454% | 2.574% | 4.781% | 3.285%
| PDM3 |  | -1.733% | 2.959% | -12.83% | -0.3508% | 2.543% | 5.075% | 0.01086% | 0.1917% | 34.84% | -5.596% | -0.5317% | 12.31% | -0.8934% | -16.45%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 2.94 | 0.08 | 0.58 | 1.16 | 0 | 0.04 | 7.97 | 1.28 | NA | 2.81 | 0.21 | 3.76
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 230b (Indicative Value 0.18 ppm) | [16 empty] | Indicative
Lu | Round | Replicate | 1.06: Lab; 1.07: Lab; 1.10: Lab; 1.11: Lab |  |  |  | 1
| No | No | 1.06: 1.06; 1.07: 1.07; 1.10: 1.19; 1.11: 1.22 |  |  |  | ppm
|  |  | 1.06: AR*MS; 1.07: AR*OES/MS; 1.10: AR*OES/MS; 1.11: AR*MS |  |  |  | 2
|  |  | 1.06: 0.2g; 1.07: 0.5g; 1.10: 0.25g; 1.11: 01g |  |  |  | 2
| 1 | 1 | 1.06: 0.18656; 1.07: 0.2; 1.10: 0.15; 1.11: < 1 |  |  |  | 1
| 1 | 2 | 1.06: 0.17549; 1.07: 0.2; 1.10: 0.15; 1.11: < 1 |  |  |  | 4
| 1 | 3 | 1.06: 0.19801; 1.07: 0.2; 1.10: 0.14; 1.11: < 1 |  |  |  | 16
| 1 | 4 | 1.06: 0.20222; 1.07: 0.2; 1.10: 0.15; 1.11: < 1 |  |  |  | 0.17702
| 1 | 5 | 1.06: 0.17782; 1.07: 0.2; 1.10: 0.14; 1.11: < 1 |  |  |  | 10
| 1 | 6 | 1.06: 0.17624; 1.07: 0.2; 1.10: 0.14; 1.11: < 1
| Mean |  | 1.06: 0.18606; 1.07: 0.2; 1.10: 0.145
| Median |  | 1.06: 0.18219; 1.07: 0.2; 1.10: 0.145
| Std Dev. |  | 1.06: 0.011664; 1.07: 3.0405e-17; 1.10: 0.0054772
| Rel.Std.Dev. |  | 1.06: 6.269%; 1.07: 1.52e-14%; 1.10: 3.777%
| PDM3 |  | 1.06: 5.106%; 1.07: 12.98%; 1.10: -18.09%
| Z-Score (Absolute) |  | 1.06: 0.17; 1.07: 0.17; 1.10: 1.18; 1.11: 7.53
| Analytical results for Mg in OREAS 230b (Certified Value 1.96 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 1.8302 | 2.1456 | 2.03 | 1.92 | 1.98 | 1.97 | 1.99 | 1.93 | 2.4358 | 1.94 | 1.86 | 1.95 | 1.95 | 2.01 |  | 1
| 1 | 2 | 1.7734 | 2.1413 | 2 | 1.91 | 2.01 | 2 | 1.99 | 1.89 | 2.3985 | 1.95 | 1.85 | 1.98 | 1.99 | 1.94 |  | 0
| 1 | 3 | 1.846 | 2.0313 | 1.99 | 1.93 | 1.99 | 1.98 | 2.01 | 1.96 | 2.2793 | 1.98 | 1.85 | 1.92 | 1.97 | 2 |  | 16
| 1 | 4 | 1.8484 | 2.141 | 2 | 1.89 | 1.96 | 2.02 | 2 | 1.98 | 2.3289 | 1.96 | 1.87 | 1.98 | 1.98 | 1.95 |  | 1.9593
| 1 | 5 | 1.834 | 2.0989 | 2.11 | 1.94 | 1.96 | 2.05 | 1.98 | 1.9 | 2.2575 | 1.97 | 1.88 | 1.95 | 1.99 | 1.96 |  | 95
| 1 | 6 | 1.7397 | 2.1154 | 1.99 | 1.91 | 1.96 | 2.03 | 2.04 | 1.91 | 2.2962 | 1.98 | 1.85 | 1.93 | 1.97 | 1.95
| Mean |  | 1.812 | 2.1123 | 2.02 | 1.9167 | 1.9767 | 2.0083 | 2.0017 | 1.9283 | 2.3327 | 1.9633 | 1.86 | 1.9517 | 1.975 | 1.9683
| Median |  | 1.8321 | 2.1282 | 2 | 1.915 | 1.97 | 2.01 | 1.995 | 1.92 | 2.3125 | 1.965 | 1.855 | 1.95 | 1.975 | 1.955
| Std Dev. |  | 0.044757 | 0.043673 | 0.046476 | 0.017512 | 0.020656 | 0.030605 | 0.02137 | 0.035449 | 0.070456 | 0.01633 | 0.012649 | 0.024833 | 0.015166 | 0.029269
| Rel.Std.Dev. |  | 2.47% | 2.068% | 2.301% | 0.9137% | 1.045% | 1.524% | 1.068% | 1.838% | 3.02% | 0.8317% | 0.6801% | 1.272% | 0.7679% | 1.487%
| PDM3 |  | -7.519% | 7.808% | 3.099% | -2.175% | 0.8878% | 2.504% | 2.164% | -1.579% | 19.06% | 0.2073% | -5.067% | -0.3882% | 0.8027% | 0.4624%
| Z-Score (Absolute) |  | 2.69 | 2.37 | 0.81 | 0.93 | 0.08 | 0.62 | 0.51 | 0.73 | 6.09 | 0.14 | 1.88 | 0.34 | 0.06 | 0.06
| Analytical results for Mn in OREAS 230b (Certified Value 0.088 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.082273 | 0.092918 | 0.0904 | 0.084 | 0.0837 | 0.0869 | 0.0878 | 0.0918 | 0.097288 | 0.0854 | 0.1025 | 0.08159 | 0.0927 | 0.0908 |  | 1
| 1 | 2 | 0.089461 | 0.092796 | 0.0914 | 0.0837 | 0.0853 | 0.0875 | 0.0881 | 0.0895 | 0.093162 | 0.086 | 0.1034 | 0.08152 | 0.0919 | 0.0893 |  | 24
| 1 | 3 | 0.086017 | 0.089739 | 0.088 | 0.084 | 0.0847 | 0.0846 | 0.0891 | 0.0932 | 0.095485 | 0.0866 | 0.1026 | 0.08118 | 0.0909 | 0.0924 |  | 16
| 1 | 4 | 0.088909 | 0.093142 | 0.0912 | 0.0836 | 0.0834 | 0.0888 | 0.0887 | 0.0893 | 0.09669 | 0.0863 | 0.1025 | 0.08163 | 0.0919 | 0.0884 |  | 0.088459
| 1 | 5 | 0.088924 | 0.091104 | 0.0931 | 0.0829 | 0.0833 | 0.0898 | 0.089 | 0.0898 | 0.093804 | 0.0868 | 0.1022 | 0.08114 | 0.0918 | 0.0898 |  | 96
| 1 | 6 | 0.0859 | 0.091842 | 0.0926 | 0.0838 | 0.0841 | 0.089 | 0.0907 | 0.0918 | 0.094205 | 0.0869 | 0.1026 | 0.0815 | 0.0924 | 0.0887
| Mean |  | 0.086914 | 0.091924 | 0.091117 | 0.083667 | 0.084083 | 0.087767 | 0.0889 | 0.0909 | 0.095106 | 0.086333 | 0.10263 | 0.081427 | 0.091933 | 0.0899
| Median |  | 0.087463 | 0.092319 | 0.0913 | 0.08375 | 0.0839 | 0.08815 | 0.08885 | 0.0908 | 0.094845 | 0.08645 | 0.10255 | 0.08151 | 0.0919 | 0.08955
| Std Dev. |  | 0.0027523 | 0.001318 | 0.0018137 | 0.00040825 | 0.00078592 | 0.0018747 | 0.0010178 | 0.00159 | 0.0016553 | 0.00056451 | 0.00040332 | 0.00021219 | 0.00061536 | 0.0014913
| Rel.Std.Dev. |  | 3.167% | 1.434% | 1.991% | 0.4879% | 0.9347% | 2.136% | 1.145% | 1.749% | 1.74% | 0.6539% | 0.393% | 0.2606% | 0.6694% | 1.659%
| PDM3 |  | -1.747% | 3.917% | 3.004% | -5.418% | -4.947% | -0.7829% | 0.4983% | 2.759% | 7.514% | -2.403% | 16.02% | -7.95% | 3.927% | 1.629%
| Z-Score (Absolute) |  | 0.66 | 0.67 | 0.46 | 1.53 | 1.42 | 0.44 | 0.13 | 0.4 | 1.52 | 0.82 | 3.53 | 2.13 | 0.68 | 0.13
| Analytical results for Mo in OREAS 230b (Certified Value 0.51 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.545 | < 1 | 0.53 | 0.47 | 0.48 | 0.64 | 0.55 | 0.51 | < 0.002 | 0.5 | 0.55 | 1.7 | 2.8 | 0.5 | 0.54 | 1
| 1 | 2 | 0.537 | < 1 | 0.5 | 0.49 | 0.53 | 0.58 | 1.41 | 0.51 | < 0.002 | 0.5 | 0.52 | 1.6 | 2.6 | 0.5 | 0.53 | 25
| 1 | 3 | 0.529 | < 1 | 0.51 | 0.46 | 0.5 | 0.52 | 0.54 | 0.52 | < 0.002 | 0.5 | 0.54 | 1.8 | 2.5 | 0.5 | 0.53 | 16
| 1 | 4 | 0.579 | < 1 | 0.54 | 0.47 | 0.48 | 0.64 | 0.52 | 0.5 | < 0.002 | 0.5 | 0.5 | 1.6 | 2.7 | 0.5 | 0.51 | 0.51429
| 1 | 5 | 0.52 | < 1 | 0.54 | 0.48 | 0.48 | 0.62 | 0.53 | 0.49 | < 0.002 | 0.5 | 0.48 | 1.9 | 2.8 | 0.5 | 0.57 | 97
| 1 | 6 |  | < 1 | 0.52 | 0.47 | 0.49 | 0.62 | 0.53 | 0.49 | < 0.002 | 0.5 | 0.49 | 1.7 | 2.7 | 0.5 | 0.51
| Mean |  | 0.542 |  | 0.52333 | 0.47333 | 0.49333 | 0.60333 | 0.68 | 0.50333 |  | 0.5 | 0.51333 | 1.7167 | 2.6833 | 0.5 | 0.53167
| Median |  | 0.537 |  | 0.525 | 0.47 | 0.485 | 0.62 | 0.535 | 0.505 |  | 0.5 | 0.51 | 1.7 | 2.7 | 0.5 | 0.53
| Std Dev. |  | 0.022672 |  | 0.01633 | 0.010328 | 0.019664 | 0.046332 | 0.35777 | 0.012111 |  | 0 | 0.028048 | 0.1169 | 0.1169 | 0 | 0.022286
| Rel.Std.Dev. |  | 4.183% |  | 3.12% | 2.182% | 3.986% | 7.679% | 52.61% | 2.406% |  | 0% | 5.464% | 6.81% | 4.357% | 0% | 4.192%
| PDM3 |  | 5.388% |  | 1.758% | -7.964% | -4.075% | 17.31% | 32.22% | -2.131% |  | -2.779% | -0.1863% | 233.8% | 421.8% | -2.779% | 3.378%
| Z-Score (Absolute) |  | 0.97 | 0.45 | 0.34 | 1.35 | 0.67 | 3.03 | 5.62 | 0.34 | 17.27 | NA | 0 | NA | NA | NA | 0.62
Results from laboratories 18, 21, 22 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 230b (Certified Value 0.108 wt.%) | [16 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.12737 | 0.10277 | 0.098 | 0.11 | 0.12 | 0.1 | 0.1 | 0.11 | 0.062932 | 0.13 | 0.1 | 0.11 | 0.11 | 0.08 |  | 1
| 1 | 2 | 0.13319 | 0.10329 | 0.094 | 0.1 | 0.12 | 0.1 | 0.11 | 0.11 | 0.060344 | 0.13 | 0.1 | 0.1 | 0.12 | 0.08 |  | 0
| 1 | 3 | 0.11783 | 0.10483 | 0.094 | 0.1 | 0.12 | 0.11 | 0.11 | 0.11 | 0.057063 | 0.12 | 0.1 | 0.11 | 0.12 | 0.08 |  | 16
| 1 | 4 | 0.13003 | 0.10528 | 0.093 | 0.1 | 0.12 | 0.1 | 0.11 | 0.11 | 0.059532 | 0.12 | 0.1 | 0.11 | 0.12 | 0.08 |  | 0.10814
| 1 | 5 | 0.12357 | 0.10093 | 0.101 | 0.11 | 0.12 | 0.11 | 0.1 | 0.1 | 0.057686 | 0.13 | 0.1 | 0.11 | 0.12 | 0.08 |  | 98
| 1 | 6 | 0.1219 | 0.10312 | 0.092 | 0.1 | 0.12 | 0.1 | 0.11 | 0.11 | 0.057726 | 0.13 | 0.1 | 0.1 | 0.12 | 0.08
| Mean |  | 0.12565 | 0.10337 | 0.095333 | 0.10333 | 0.12 | 0.10333 | 0.10667 | 0.10833 | 0.059214 | 0.12667 | 0.1 | 0.10667 | 0.11833 | 0.08
| Median |  | 0.12547 | 0.1032 | 0.094 | 0.1 | 0.12 | 0.1 | 0.11 | 0.11 | 0.058629 | 0.13 | 0.1 | 0.11 | 0.12 | 0.08
| Std Dev. |  | 0.0056318 | 0.0015572 | 0.0034448 | 0.005164 | 0 | 0.005164 | 0.005164 | 0.0040825 | 0.0022079 | 0.005164 | 1.5202e-17 | 0.005164 | 0.0040825 | 0
| Rel.Std.Dev. |  | 4.482% | 1.506% | 3.613% | 4.997% | 0% | 4.997% | 4.841% | 3.768% | 3.729% | 4.077% | 1.52e-14% | 4.841% | 3.45% | 0%
| PDM3 |  | 16.19% | -4.412% | -11.84% | -4.445% | 10.97% | -4.445% | -1.363% | 0.1786% | -45.24% | 17.13% | -7.527% | -1.363% | 9.426% | -26.02%
| Z-Score (Absolute) |  | 1.13 | 0.2 | 0.67 | 0.2 | 0.79 | 0.2 | 0 | 0.1 | 2.82 | 1.19 | NA | 0 | 0.69 | 1.59
Results from laboratory 19 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb in OREAS 230b (Indicative Value 0.16 ppm) | [16 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 0.14459 | 0.2 | 0.13 | 0.11 | < 0.05 | 0.23 | 0.16 | 0.1533 | 0.3 | 0.11 | 3.7 | 0.3 | 0.09 |  |  | 1
| 1 | 2 | 0.13738 | 0.1 | 0.13 | 0.11 | < 0.05 | 0.23 | 0.16 | 0.22453 | 0.3 | 0.11 | 3.7 | 0.3 | 0.09 |  |  | 5
| 1 | 3 | 0.13898 | 0.2 | 0.13 | 0.11 | 0.16 | 0.23 | 0.18 | 0.12823 | 0.3 | 0.1 | 3.7 | 0.3 | 0.09 |  |  | 16
| 1 | 4 | 0.12674 | 0.1 | 0.12 | 0.1 | 0.06 | 0.21 | 0.16 | 0.18478 | 0.2 | 0.11 | 3.7 | 0.4 | 0.1 |  |  | 0.16266
| 1 | 5 | 0.14542 | 0.1 | 0.13 | 0.1 | 0.07 | 0.23 | 0.15 | 0.17905 | 0.3 | 0.1 | 3.7 | 0.3 | 0.09 |  |  | 11
| 1 | 6 | 0.14557 | 0.2 | 0.14 | 0.1 | < 0.05 | 0.23 | 0.17 | 0.19268 | 0.2 | 0.1 | 3.8 | 0.3 | 0.09
| Mean |  | 0.13978 | 0.15 | 0.13 | 0.105 | 0.096667 | 0.22667 | 0.16333 | 0.17709 | 0.26667 | 0.105 | 3.7167 | 0.31667 | 0.091667
| Median |  | 0.14179 | 0.15 | 0.13 | 0.105 | 0.07 | 0.23 | 0.16 | 0.18191 | 0.3 | 0.105 | 3.7 | 0.3 | 0.09
| Std Dev. |  | 0.007282 | 0.054772 | 0.0063246 | 0.0054772 | 0.055076 | 0.008165 | 0.010328 | 0.033194 | 0.05164 | 0.0054772 | 0.040825 | 0.040825 | 0.0040825
| Rel.Std.Dev. |  | 5.21% | 36.51% | 4.865% | 5.216% | 56.97% | 3.602% | 6.323% | 18.74% | 19.36% | 5.216% | 1.098% | 12.89% | 4.454%
| PDM3 |  | -14.06% | -7.781% | -20.08% | -35.45% | -40.57% | 39.35% | 0.4162% | 8.877% | 63.94% | -35.45% | 2185% | 94.68% | -43.64%
| Z-Score (Absolute) |  | 0.15 | 0 | 0.3 | 0.67 | 1.34 | 1.15 | 0.2 | 0.41 | 1.75 | 0.67 | 53.45 | 2.5 | 0.87
| Analytical results for Nd in OREAS 230b (Indicative Value 4.74 ppm) | [16 empty] | Indicative
Nd | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 4.7986; 1.06: 4.93 | 1
| 1 | 2 | 1.05: 4.6845; 1.06: 4.79 | 6
| 1 | 3 | 1.05: 4.4469; 1.06: 4.8 | 16
| 1 | 4 | 1.05: 4.5567; 1.06: 4.89 | 4.7443
| 1 | 5 | 1.05: 4.4829; 1.06: 4.81 | 12
| 1 | 6 | 1.05: 4.7621; 1.06: 4.98
| Mean |  | 1.05: 4.6219; 1.06: 4.8667
| Median |  | 1.05: 4.6206; 1.06: 4.85
| Std Dev. |  | 1.05: 0.14764; 1.06: 0.078655
| Rel.Std.Dev. |  | 1.05: 3.194%; 1.06: 1.616%
| PDM3 |  | 1.05: -2.579%; 1.06: 2.579%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Ni in OREAS 230b (Certified Value 44.4 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 42.009 | 46.3058 | 46.7 | 43.1 | 46.4 | 42.1 | 46.2 | 42.1 | 48.8579 | 44 | 48.7 | 37.6 | 40.4 | 44.6 |  | 1
| 1 | 2 | 44.58 | 45.529 | 42.6 | 43.2 | 45.4 | 42.6 | 52.6 | 40.6 | 48.0878 | 45 | 47.8 | 37.7 | 40.4 | 42.6 |  | 28
| 1 | 3 | 41.932 | 44.9317 | 42.6 | 42.5 | 46.6 | 43.3 | 48.6 | 42.8 | 45.9111 | 45 | 48.8 | 36.9 | 40.7 | 44.4 |  | 16
| 1 | 4 | 43.407 | 47.2559 | 44 | 42.9 | 45.8 | 42.4 | 45.6 | 42.9 | 46.5967 | 44 | 47.6 | 37.4 | 40.3 | 42.5 |  | 44.4146
| 1 | 5 | 45.162 | 45.1623 | 45.9 | 43.5 | 46.4 | 42.9 | 47.3 | 39.3 | 45.144 | 44 | 47.7 | 37.1 | 39.9 | 43.1 |  | 99
| 1 | 6 | 47 | 45.5747 | 45 | 43.4 | 45.7 | 42.9 | 48.3 | 41.1 | 45.9921 | 44 | 46.6 | 37.7 | 40 | 42.9
| Mean |  | 44.015 | 45.7932 | 44.4667 | 43.1 | 46.05 | 42.7 | 48.1 | 41.4667 | 46.7649 | 44.3333 | 47.8667 | 37.4 | 40.2833 | 43.35
| Median |  | 43.9935 | 45.5519 | 44.5 | 43.15 | 46.1 | 42.75 | 47.8 | 41.6 | 46.2944 | 44 | 47.75 | 37.5 | 40.35 | 43
| Std Dev. |  | 1.9636 | 0.85559 | 1.7037 | 0.36332 | 0.48062 | 0.42426 | 2.4916 | 1.401 | 1.4221 | 0.5164 | 0.80911 | 0.33466 | 0.29269 | 0.91815
| Rel.Std.Dev. |  | 4.461% | 1.868% | 3.831% | 0.843% | 1.044% | 0.9936% | 5.18% | 3.379% | 3.041% | 1.165% | 1.69% | 0.8948% | 0.7266% | 2.118%
| PDM3 |  | -0.8997% | 3.104% | 0.1172% | -2.96% | 3.682% | -3.86% | 8.298% | -6.637% | 5.292% | -0.183% | 7.772% | -15.79% | -9.302% | -2.397%
| Z-Score (Absolute) |  | 0.06 | 0.62 | 0.11 | 0.41 | 0.72 | 0.57 | 1.51 | 1.04 | 1 | 0.06 | 1.42 | 2.61 | 1.5 | 0.32
| Analytical results for P in OREAS 230b (Certified Value 0.048 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.044725 | 0.049291 | 0.046 | 0.046 | 0.047 | 0.051 | 0.049 | 0.045 | 0.0475 | 0.0475 | 0.045 | 0.054 | 0.0481 |  |  | 1
| 1 | 2 | 0.044459 | 0.048514 | 0.044 | 0.046 | 0.048 | 0.052 | 0.049 | 0.045 | 0.0482 | 0.0502 | 0.045 | 0.052 | 0.0469 |  |  | 0
| 1 | 3 | 0.045323 | 0.047891 | 0.044 | 0.046 | 0.047 | 0.049 | 0.049 | 0.046 | 0.0483 | 0.0489 | 0.044 | 0.053 | 0.0484 |  |  | 16
| 1 | 4 | 0.044776 | 0.049545 | 0.044 | 0.046 | 0.046 | 0.052 | 0.05 | 0.046 | 0.0482 | 0.0494 | 0.044 | 0.051 | 0.0471 |  |  | 0.047581
| 1 | 5 | 0.045284 | 0.048533 | 0.044 | 0.047 | 0.047 | 0.053 | 0.049 | 0.044 | 0.0481 | 0.0501 | 0.043 | 0.05 | 0.0487 |  |  | 100
| 1 | 6 | 0.0437 | 0.048206 | 0.044 | 0.046 | 0.047 | 0.053 | 0.051 | 0.045 | 0.0482 | 0.0488 | 0.044 | 0.052 | 0.0485
| Mean |  | 0.044711 | 0.048663 | 0.044333 | 0.046167 | 0.047 | 0.051667 | 0.0495 | 0.045167 | 0.048083 | 0.04915 | 0.044167 | 0.052 | 0.04795
| Median |  | 0.044751 | 0.048524 | 0.044 | 0.046 | 0.047 | 0.052 | 0.049 | 0.045 | 0.0482 | 0.04915 | 0.044 | 0.052 | 0.04825
| Std Dev. |  | 0.00059885 | 0.00063492 | 0.0008165 | 0.00040825 | 0.00063246 | 0.0015055 | 0.00083666 | 0.00075277 | 0.00029269 | 0.0009975 | 0.00075277 | 0.0014142 | 0.00076354
| Rel.Std.Dev. |  | 1.339% | 1.305% | 1.842% | 0.8843% | 1.346% | 2.914% | 1.69% | 1.667% | 0.6087% | 2.029% | 1.704% | 2.72% | 1.592%
| PDM3 |  | -6.032% | 2.274% | -6.826% | -2.973% | -1.222% | 8.586% | 4.032% | -5.075% | 1.055% | 3.297% | -7.177% | 9.286% | 0.7747%
| Z-Score (Absolute) |  | 1.22 | 0.27 | 1.37 | 0.67 | 0.36 | 1.41 | 0.59 | 1.05 | 0.05 | 0.45 | 1.43 | 1.53 | 0
| Analytical results for Pb in OREAS 230b (Certified Value 8.41 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 9.265 | 8.9994 | 8.7 | 8 | 7.7 | 8.5 | 8.9 | 8.2 | 4.6663 | 8.5 | 8.1 | 8 | 7.6 | 6.4 | 8.6 | 1
| 1 | 2 | 9.458 | 9.2554 | 8.4 | 8.1 | 7.9 | 8.2 | 15.6 | 7.9 | 4.6311 | 8.5 | 8.2 | 7 | 7.9 | 6.2 | 8.7 | 29
| 1 | 3 | 8.825 | 8.9203 | 8.3 | 8.4 | 7.7 | 8.3 | 9 | 8.2 | 6.0376 | 8.6 | 8.1 | 7 | 7.6 | 6.4 | 8.8 | 16
| 1 | 4 | 9.066 | 8.6942 | 8.8 | 8.1 | 7.6 | 8.6 | 8.9 | 8.3 | 4.8772 | 8.5 | 8.1 | 6 | 7.6 | 6.4 | 9.7 | 8.4112
| 1 | 5 | 8.909 | 8.6427 | 8.8 | 7.9 | 7.5 | 8.7 | 8.8 | 8 | 4.578 | 8.5 | 8.1 | 11 | 8 | 6.7 | 8.8 | 101
| 1 | 6 | 9 | 9.0527 | 8.6 | 8 | 7.7 | 8.7 | 8.9 | 8 | 5.5199 | 8.5 | 7.9 | 7 | 7.7 | 6.7 | 8.7
| Mean |  | 9.0872 | 8.9275 | 8.6 | 8.0833 | 7.6833 | 8.5 | 10.0167 | 8.1 | 5.0517 | 8.5167 | 8.0833 | 7.6667 | 7.7333 | 6.4667 | 8.8833
| Median |  | 9.033 | 8.9599 | 8.65 | 8.05 | 7.7 | 8.55 | 8.9 | 8.1 | 4.7717 | 8.5 | 8.1 | 7 | 7.65 | 6.4 | 8.75
| Std Dev. |  | 0.23565 | 0.22977 | 0.20976 | 0.17224 | 0.13292 | 0.20976 | 2.736 | 0.15492 | 0.59524 | 0.040825 | 0.098319 | 1.7512 | 0.17512 | 0.19664 | 0.40702
| Rel.Std.Dev. |  | 2.593% | 2.574% | 2.439% | 2.131% | 1.73% | 2.468% | 27.31% | 1.913% | 11.78% | 0.4794% | 1.216% | 22.84% | 2.264% | 3.041% | 4.582%
| PDM3 |  | 8.036% | 6.138% | 2.244% | -3.898% | -8.654% | 1.055% | 19.09% | -3.7% | -39.94% | 1.254% | -3.898% | -8.852% | -8.059% | -23.12% | 5.613%
| Z-Score (Absolute) |  | 0.92 | 0.74 | 0.35 | 0.25 | 0.72 | 0.23 | 2.01 | 0.23 | 3.81 | 0.25 | 0.25 | NA | 0.66 | 2.15 | 0.68
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 230b (Indicative Value 0.99 ppm) | [16 empty] | Indicative
Pr | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 0.97049; 1.06: 1 | 1
| 1 | 2 | 1.05: 0.93052; 1.06: 1 | 7
| 1 | 3 | 1.05: 0.95967; 1.06: 1 | 16
| 1 | 4 | 1.05: 0.97438; 1.06: 1 | 0.98523
| 1 | 5 | 1.05: 0.92581; 1.06: 1.1 | 13
| 1 | 6 | 1.05: 0.96192; 1.06: 1
| Mean |  | 1.05: 0.9538; 1.06: 1.0167
| Median |  | 1.05: 0.96079; 1.06: 1
| Std Dev. |  | 1.05: 0.020627; 1.06: 0.040825
| Rel.Std.Dev. |  | 1.05: 2.163%; 1.06: 4.016%
| PDM3 |  | 1.05: -3.191%; 1.06: 3.191%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Rb in OREAS 230b (Certified Value 2.24 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 2.545 | 2.1482 | 2 | 2.1 | 2.1 | 2.3 | 2.4 | 2.3 | 2.38 | 2.2 | 4.8 | 2.4 | 2 |  |  | 1
| 1 | 2 | 2.714 | 2.1937 | 1.9 | 2.1 | 2.2 | 2.3 | 2.4 | 2.2 | 2.32 | 2.2 | 4.8 | 2.4 | 2 |  |  | 31
| 1 | 3 | 2.455 | 2.2045 | 1.8 | 2.1 | 2.1 | 2.4 | 2.4 | 2.3 | 2.36 | 2.2 | 4.9 | 2.4 | 2 |  |  | 16
| 1 | 4 | 2.588 | 2.131 | 1.9 | 2.1 | 2.1 | 2.3 | 2.3 | 2.3 | 2.33 | 2.2 | 4.7 | 2.4 | 2 |  |  | 2.2428
| 1 | 5 | 2.434 | 2.1713 | 2 | 2.1 | 2 | 2.3 | 2.3 | 2.1 | 2.4 | 2.2 | 4.8 | 2.4 | 2 |  |  | 102
| 1 | 6 | 3 | 2.0529 | 1.9 | 2.2 | 2.1 | 2.3 | 2.3 | 2.2 | 2.35 | 2.1 | 4.7 | 2.4 | 2
| Mean |  | 2.6227 | 2.1503 | 1.9167 | 2.1167 | 2.1 | 2.3167 | 2.35 | 2.2333 | 2.3567 | 2.1833 | 4.7833 | 2.4 | 2
| Median |  | 2.5665 | 2.1598 | 1.9 | 2.1 | 2.1 | 2.3 | 2.35 | 2.25 | 2.355 | 2.2 | 4.8 | 2.4 | 2
| Std Dev. |  | 0.21055 | 0.054988 | 0.075277 | 0.040825 | 0.063246 | 0.040825 | 0.054772 | 0.08165 | 0.030111 | 0.040825 | 0.075277 | 0 | 0
| Rel.Std.Dev. |  | 8.028% | 2.557% | 3.928% | 1.929% | 3.012% | 1.762% | 2.331% | 3.656% | 1.278% | 1.87% | 1.574% | 0% | 0%
| PDM3 |  | 16.94% | -4.126% | -14.54% | -5.624% | -6.367% | 3.294% | 4.78% | -0.4221% | 5.077% | -2.651% | 113.3% | 7.009% | -10.83%
| Z-Score (Absolute) |  | 1.88 | 0.67 | 1.94 | 0.86 | 0.95 | 0.23 | 0.41 | 0.23 | 0.44 | 0.5 | 13.55 | 0.68 | NA
Results from laboratory 25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Re in OREAS 230b (Indicative Value 0.002 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | < 0.01 | 0.002 | 0.002 | 0.001 | 0.001 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.5 | < 0.005 | 0.003 |  |  |  | 1
| 1 | 2 | < 0.01 | 0.002 | 0.002 | 0.002 | 0.001 | 0.002 | 0.002 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.003 |  |  |  | 8
| 1 | 3 | < 0.01 | 0.002 | 0.002 | 0.002 | 0.003 | 0.002 | 0.002 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.003 |  |  |  | 16
| 1 | 4 | < 0.01 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.003 |  |  |  | 0.002125
| 1 | 5 | < 0.01 | 0.002 | 0.002 | 0.002 | 0.003 | 0.002 | 0.002 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.003 |  |  |  | 14
| 1 | 6 | < 0.01 | 0.002 | 0.002 | 0.002 | 0.003 | 0.002 | 0.002 | < 0.05 | < 0.002 | < 0.5 | < 0.005 | 0.003
| Mean |  | 1.06: 0.002; 1.07: 0.002; 1.10: 0.0018333; 1.11: 0.0021667; 1.12: 0.002; 1.13: 0.002; 1.17: 0.002; 1.21: 0.003
| Median |  | 1.06: 0.002; 1.07: 0.002; 1.10: 0.002; 1.11: 0.0025; 1.12: 0.002; 1.13: 0.002; 1.17: 0.002; 1.21: 0.003
| Std Dev. |  | 1.06: 0; 1.07: 0; 1.10: 0.00040825; 1.11: 0.00098319; 1.12: 0; 1.13: 0; 1.21: 4.7507e-19
| Rel.Std.Dev. |  | 1.06: 0%; 1.07: 0%; 1.10: 22.27%; 1.11: 45.38%; 1.12: 0%; 1.13: 0%; 1.21: 1.584e-14%
| PDM3 |  | 1.06: -5.882%; 1.07: -5.882%; 1.10: -13.73%; 1.11: 1.961%; 1.12: -5.882%; 1.13: -5.882%; 1.17: -5.882%; 1.21: 41.18%
| Z-Score (Absolute) |  | 5.9 | 0.17 | 0.17 | 0.51 | 0.17 | 0.17 | 0.17 | 46.36 | 1.85 | 501.52 | 0.84 | 1.85
| Analytical results for S in OREAS 230b (Certified Value 0.266 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.20937 | 0.26976 | 0.243 | 0.27 | 0.27 | 0.27 | 0.27 | 0.26 | 0.26 | 0.3 | 0.2716 | 0.15 | 0.27 | 0.3 |  | 1
| 1 | 2 | 0.20914 | 0.27364 | 0.238 | 0.26 | 0.27 | 0.27 | 0.28 | 0.26 | 0.26 | 0.29 | 0.2587 | 0.15 | 0.27 | 0.29 |  | 13
| 1 | 3 | 0.20051 | 0.27368 | 0.244 | 0.26 | 0.27 | 0.26 | 0.27 | 0.27 | 0.26 | 0.3 | 0.2648 | 0.15 | 0.27 | 0.29 |  | 16
| 1 | 4 | 0.22113 | 0.27202 | 0.235 | 0.26 | 0.26 | 0.28 | 0.27 | 0.25 | 0.26 | 0.29 | 0.2616 | 0.15 | 0.27 | 0.29 |  | 0.26634
| 1 | 5 | 0.21855 | 0.26303 | 0.243 | 0.27 | 0.26 | 0.28 | 0.27 | 0.25 | 0.26 | 0.29 | 0.2633 | 0.15 | 0.27 | 0.29 |  | 103
| 1 | 6 | 0.2008 | 0.26953 | 0.247 | 0.26 | 0.26 | 0.28 | 0.28 | 0.26 | 0.26 | 0.3 | 0.257 | 0.15 | 0.27 | 0.29
| Mean |  | 0.20992 | 0.27028 | 0.24167 | 0.26333 | 0.265 | 0.27333 | 0.27333 | 0.25833 | 0.26 | 0.295 | 0.26283 | 0.15 | 0.27 | 0.29167
| Median |  | 0.20925 | 0.27089 | 0.243 | 0.26 | 0.265 | 0.275 | 0.27 | 0.26 | 0.26 | 0.295 | 0.26245 | 0.15 | 0.27 | 0.29
| Std Dev. |  | 0.0086348 | 0.0039805 | 0.0043665 | 0.005164 | 0.0054772 | 0.008165 | 0.005164 | 0.0075277 | 0 | 0.0054772 | 0.0051686 | 0 | 0 | 0.0040825
| Rel.Std.Dev. |  | 4.113% | 1.473% | 1.807% | 1.961% | 2.067% | 2.987% | 1.889% | 2.914% | 0% | 1.857% | 1.967% | 0% | 0% | 1.4%
| PDM3 |  | -21.19% | 1.478% | -9.265% | -1.13% | -0.5044% | 2.624% | 2.624% | -3.007% | -2.382% | 10.76% | -1.318% | -43.68% | 1.373% | 9.508%
| Z-Score (Absolute) |  | 4.79 | 0.54 | 1.99 | 0.07 | 0.07 | 0.81 | 0.81 | 0.51 | 0.37 | 2.72 | 0.12 | 10.08 | 0.51 | 2.43
| Analytical results for Sb in OREAS 230b (Certified Value 0.52 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 0.516 | < 2.5 | 0.4 | 0.56 | 0.58 | 0.45 | 0.33 | 0.56 | 0.59879 | 0.54 | 0.52 | 4.07 | 3.4 | 0.85 | 2.73 | 1
| 1 | 2 | 0.526 | < 2.5 | 0.39 | 0.57 | 0.6 | 0.46 | 0.85 | 0.62 | 0.56563 | 0.33 | 0.55 | 1.5 | 3.4 | 0.88 | 2.6 | 33
| 1 | 3 | 0.453 | < 2.5 | 0.41 | 0.7 | 0.63 | 0.59 | 0.39 | 0.61 | 0.63575 | 0.27 | 0.56 | 2.72 | 3.1 | 0.87 | 2.51 | 16
| 1 | 4 | 0.607 | < 2.5 | 0.38 | 0.58 | 0.56 | 0.49 | 0.32 | 0.64 | 0.43285 | 0.22 | 0.54 | 0.96 | 3.1 | 0.86 | 2.82 | 0.51635
| 1 | 5 | 0.537 | < 2.5 | 0.42 | 0.59 | 0.56 | 0.51 | 0.37 | 0.54 | 0.63876 | 0.47 | 0.53 | 5.38 | 3.6 | 0.86 | 2.96 | 104
| 1 | 6 |  | < 2.5 | 0.8 | 0.62 | 0.57 | 0.45 | 0.31 | 0.56 | 0.65237 | 0.26 | 0.53 | 0.82 | 3 | 0.89 | 2.91
| Mean |  | 0.5278 |  | 0.46667 | 0.60333 | 0.58333 | 0.49167 | 0.42833 | 0.58833 | 0.58736 | 0.34833 | 0.53833 | 2.575 | 3.2667 | 0.86833 | 2.755
| Median |  | 0.526 |  | 0.405 | 0.585 | 0.575 | 0.475 | 0.35 | 0.585 | 0.61727 | 0.3 | 0.535 | 2.11 | 3.25 | 0.865 | 2.775
| Std Dev. |  | 0.054988 |  | 0.16391 | 0.05164 | 0.027325 | 0.053821 | 0.20885 | 0.040208 | 0.08208 | 0.12828 | 0.01472 | 1.8422 | 0.23381 | 0.01472 | 0.17604
| Rel.Std.Dev. |  | 10.42% |  | 35.12% | 8.559% | 4.684% | 10.95% | 48.76% | 6.834% | 13.97% | 36.83% | 2.734% | 71.54% | 7.157% | 1.695% | 6.39%
| PDM3 |  | 2.218% |  | -9.622% | 16.85% | 12.97% | -4.78% | -17.05% | 13.94% | 13.75% | -32.54% | 4.258% | 398.7% | 532.6% | 68.17% | 433.6%
| Z-Score (Absolute) |  | 0.37 | 4.22 | 0.75 | 0.11 | 0.01 | 0.59 | 1 | 0.02 | 0.01 | 1.51 | 0.3 | 12.64 | NA | 1.8 | 13.78
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 230b (Certified Value 6.01 ppm) | [16 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 3.348 | 6.3054 | 8.3 | 5.6 | 5.7 | 6.2 | 6.8 | 7.3 | 7 | 4.9 | 6.1 | 7.9 | 3.8 |  |  | 1
| 1 | 2 | 3.571 | 6.254 | 7.1 | 5.8 | 5.7 | 6.3 | 7 | 7.1 | 7 | 4.9 | 6 | 7.7 | 3.7 |  |  | 34
| 1 | 3 | 3.902 | 6.211 | 7.9 | 5.7 | 5.8 | 6.4 | 7.5 | 7.3 | 7 | 4.7 | 5.8 | 7.9 | 3.7 |  |  | 16
| 1 | 4 | 3.932 | 6.3379 | 7.7 | 5.6 | 5.6 | 6.4 | 6.7 | 7.1 | 7 | 5 | 5.9 | 7.8 | 3.8 |  |  | 6.0064
| 1 | 5 | 3.745 | 6.1539 | 7.9 | 5.8 | 5.7 | 6.6 | 7 | 7 | 7 | 4.8 | 5.9 | 7.7 | 3.8 |  |  | 105
| 1 | 6 | 3.836 | 6.3625 | 7.8 | 5.8 | 5.8 | 6.6 | 7 | 7 | 7 | 4.7 | 6 | 7.7 | 3.7
| Mean |  | 3.7223 | 6.2708 | 7.7833 | 5.7167 | 5.7167 | 6.4167 | 7 | 7.1333 | 7 | 4.8333 | 5.95 | 7.7833 | 3.75
| Median |  | 3.7905 | 6.2797 | 7.85 | 5.75 | 5.7 | 6.4 | 7 | 7.1 | 7 | 4.85 | 5.95 | 7.75 | 3.75
| Std Dev. |  | 0.22481 | 0.079528 | 0.392 | 0.098319 | 0.075277 | 0.16021 | 0.27568 | 0.13663 | 0 | 0.12111 | 0.10488 | 0.098319 | 0.054772
| Rel.Std.Dev. |  | 6.04% | 1.268% | 5.036% | 1.72% | 1.317% | 2.497% | 3.938% | 1.915% | 0% | 2.506% | 1.763% | 1.263% | 1.461%
| PDM3 |  | -38.03% | 4.402% | 29.58% | -4.823% | -4.823% | 6.831% | 16.54% | 18.76% | 16.54% | -19.53% | -0.9385% | 29.58% | -37.57%
| Z-Score (Absolute) |  | 1.68 | 0.11 | 1.18 | 0.28 | 0.28 | 0.22 | 0.63 | 0.72 | NA | 0.9 | 0.11 | 1.18 | 1.66
Results from laboratory 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 230b (Certified Value 0.58 ppm) | [16 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 0.588 | < 5 | 0.6 | 0.5 | 0.6 | 0.2 | 0.6 | 0.6 | < 1 | < 1 | 0.63 | < 0.5 | 1.4 | 0.4 |  | 1
| 1 | 2 | 0.356 | < 5 | 0.8 | 0.6 | 0.6 | 0.8 | 0.7 | 0.6 | < 1 | < 1 | 0.68 | < 0.5 | 1.4 | 0.4 |  | 0
| 1 | 3 | 0.534 | < 5 | 0.7 | 0.5 | 0.5 | 0.6 | 0.7 | 0.6 | < 1 | < 1 | 0.58 | < 0.5 | 1.4 | 0.4 |  | 16
| 1 | 4 | 0.534 | < 5 | 0.6 | 0.4 | 0.6 | 0.5 | 0.6 | 0.5 | < 1 | < 1 | 0.59 | < 0.5 | 1.7 | 0.4 |  | 0.58389
| 1 | 5 | 0.569 | < 5 | 0.7 | 0.6 | 0.6 | 0.6 | 0.7 | 0.6 | < 1 | < 1 | 0.63 | < 0.5 | 1.2 | 0.4 |  | 106
| 1 | 6 | 0.45 | < 5 | 0.7 | 0.5 | 0.6 | 0.8 | 0.7 | 0.6 | < 1 | < 1 | 0.65 | < 0.5 | 1.4 | 0.4
| Mean |  | 0.50517 |  | 0.68333 | 0.51667 | 0.58333 | 0.58333 | 0.66667 | 0.58333 |  |  | 0.62667 |  | 1.4167 | 0.4
| Median |  | 0.534 |  | 0.7 | 0.5 | 0.6 | 0.6 | 0.7 | 0.6 |  |  | 0.63 |  | 1.4 | 0.4
| Std Dev. |  | 0.087055 |  | 0.075277 | 0.075277 | 0.040825 | 0.22286 | 0.05164 | 0.040825 |  |  | 0.037238 |  | 0.16021 | 6.0809e-17
| Rel.Std.Dev. |  | 17.23% |  | 11.02% | 14.57% | 6.999% | 38.2% | 7.746% | 6.999% |  |  | 5.942% |  | 11.31% | 1.52e-14%
| PDM3 |  | -13.48% |  | 17.03% | -11.51% | -0.09515% | -0.09515% | 14.18% | -0.09515% |  |  | 7.326% |  | 142.6% | -31.49%
| Z-Score (Absolute) |  | 0.63 | 15.51 | 0.81 | 0.54 | 0 | 0 | 0.67 | 0 | 0.67 | 0.67 | 0.35 | 2.7 | 6.74 | 1.48
| Analytical results for Sm in OREAS 230b (Indicative Value 1.51 ppm) | [16 empty] | Indicative
Sm | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 1.4361; 1.06: 1.5 | 1
| 1 | 2 | 1.05: 1.3726; 1.06: 1.8 | 35
| 1 | 3 | 1.05: 1.3301; 1.06: 1.5 | 16
| 1 | 4 | 1.05: 1.379; 1.06: 1.6 | 1.5122
| 1 | 5 | 1.05: 1.4612; 1.06: 1.8 | 9
| 1 | 6 | 1.05: 1.3674; 1.06: 1.6
| Mean |  | 1.05: 1.3911; 1.06: 1.6333
| Median |  | 1.05: 1.3758; 1.06: 1.6
| Std Dev. |  | 1.05: 0.048393; 1.06: 0.13663
| Rel.Std.Dev. |  | 1.05: 3.479%; 1.06: 8.365%
| PDM3 |  | 1.05: -8.011%; 1.06: 8.011%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for Sn in OREAS 230b (Certified Value 0.46 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 0.4 | 0.58072 | 0.54 | 0.4 | 0.4 | 0.5 | 0.7 | 0.5 | 0.28898 | 0.6 | 0.5 | 2.5 | 0.5 | 0.3 |  | 1
| 1 | 2 | 0.371 | 0.60991 | 0.5 | 0.4 | 0.4 | 0.5 | 0.7 | 0.5 | 0.25105 | 0.6 | 0.4 | 2.8 | 0.5 | 0.3 |  | 19
| 1 | 3 | 0.328 | 0.59227 | 0.49 | 0.4 | 0.4 | 0.4 | 0.7 | 0.5 | 0.28566 | 0.6 | 0.4 | 2.3 | 0.5 | 0.3 |  | 16
| 1 | 4 | 0.395 | 0.55379 | 0.5 | 0.4 | 0.5 | 0.5 | 0.7 | 0.5 | 0.23988 | 0.6 | 0.4 | 3.1 | 0.5 | 0.3 |  | 0.46344
| 1 | 5 | 0.355 | 0.56567 | 0.55 | 0.4 | 0.4 | 0.5 | 0.7 | 0.5 | 0.27782 | 0.6 | 0.4 | 2.8 | 0.5 | 0.3 |  | 107
| 1 | 6 | 0.314 | 0.58967 | 0.49 | 0.4 | 0.4 | 0.5 | 0.6 | 0.5 | 0.28007 | 0.6 | 0.4 | 2.7 | 0.5 | 0.3
| Mean |  | 0.3605 | 0.58201 | 0.51167 | 0.4 | 0.41667 | 0.48333 | 0.68333 | 0.5 | 0.27057 | 0.6 | 0.41667 | 2.7 | 0.5 | 0.3
| Median |  | 0.363 | 0.58519 | 0.5 | 0.4 | 0.4 | 0.5 | 0.7 | 0.5 | 0.27895 | 0.6 | 0.4 | 2.75 | 0.5 | 0.3
| Std Dev. |  | 0.034956 | 0.020038 | 0.026394 | 6.0809e-17 | 0.040825 | 0.040825 | 0.040825 | 0 | 0.020163 | 0 | 0.040825 | 0.27568 | 0 | 0
| Rel.Std.Dev. |  | 9.696% | 3.443% | 5.159% | 1.52e-14% | 9.798% | 8.447% | 5.974% | 0% | 7.452% | 0% | 9.798% | 10.21% | 0% | 0%
| PDM3 |  | -22.21% | 25.58% | 10.41% | -13.69% | -10.09% | 4.292% | 47.45% | 7.888% | -41.62% | 29.47% | -10.09% | 482.6% | 7.888% | -35.27%
| Z-Score (Absolute) |  | 0.97 | 0.67 | 0.15 | 0.68 | 0.56 | 0.06 | 1.42 | 0.06 | 1.64 | 0.8 | 0.56 | 16.36 | 0.06 | 1.42
| Analytical results for Sr in OREAS 230b (Certified Value 23.5 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 1
| 1 | 1 | 18.493 | 25.071 | 37.3 | 24.5 | 22.3 | 23.9 | 28.3 | 27.3 | 22.3867 | 36.4 | 19.2 | 27.8 | 24 | 17.6 |  | 1
| 1 | 2 | 18.69 | 24.8527 | 34.6 | 24.8 | 22.8 | 24.3 | 28.6 | 27.2 | 22.3914 | 35.6 | 19.3 | 27.3 | 25 | 17.1 |  | 20
| 1 | 3 | 18.696 | 25.1868 | 35.4 | 25.7 | 22.4 | 22.9 | 29 | 28 | 22.6115 | 37 | 18.7 | 26.5 | 24 | 17.1 |  | 16
| 1 | 4 | 18.679 | 25.2071 | 34.6 | 24.6 | 21.9 | 24.8 | 28.6 | 26.4 | 22.8206 | 35.8 | 18.9 | 27.7 | 25 | 17.3 |  | 23.487
| 1 | 5 | 20.643 | 23.7181 | 34.4 | 24.9 | 21.8 | 25.5 | 28.6 | 26.9 | 21.8036 | 35.9 | 19.1 | 27.4 | 25 | 17.6 |  | 108
| 1 | 6 | 19.599 | 25.0407 | 36.8 | 25.2 | 22.4 | 25.1 | 28.7 | 27 | 22.3342 | 36.4 | 18.5 | 27.3 | 24 | 17.5
| Mean |  | 19.1333 | 24.8461 | 35.5167 | 24.95 | 22.2667 | 24.4167 | 28.6333 | 27.1333 | 22.3913 | 36.1833 | 18.95 | 27.3333 | 24.5 | 17.3667
| Median |  | 18.693 | 25.0559 | 35 | 24.85 | 22.35 | 24.55 | 28.6 | 27.1 | 22.389 | 36.15 | 19 | 27.35 | 24.5 | 17.4
| Std Dev. |  | 0.83668 | 0.56696 | 1.2465 | 0.44159 | 0.36697 | 0.9347 | 0.22509 | 0.52789 | 0.34101 | 0.51543 | 0.30822 | 0.45898 | 0.54772 | 0.23381
| Rel.Std.Dev. |  | 4.373% | 2.282% | 3.51% | 1.77% | 1.648% | 3.828% | 0.7861% | 1.946% | 1.523% | 1.424% | 1.626% | 1.679% | 2.236% | 1.346%
| PDM3 |  | -18.54% | 5.786% | 51.22% | 6.229% | -5.196% | 3.958% | 21.91% | 15.52% | -4.665% | 54.06% | -19.32% | 16.38% | 4.313% | -26.06%
| Z-Score (Absolute) |  | 1.46 | 0.05 | 2.86 | 0.07 | 0.63 | 0.07 | 1.04 | 0.65 | 0.6 | 3.03 | 1.51 | 0.7 | 0.05 | 1.92
| Analytical results for Ta in OREAS 230b (Certified Value < 0.01 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | < 0.01 | < 0.05 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | 0.024495 | < 0.05 | < 0.05 | < 0.5 | < 0.1 | < 0.01 |  |  | 1
| 1 | 2 | < 0.01 | < 0.05 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | 0.027229 | < 0.05 | < 0.05 | < 0.5 | < 0.1 | < 0.01 |  |  | 21
| 1 | 3 | < 0.01 | < 0.05 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | 0.038747 | < 0.05 | < 0.05 | < 0.5 | < 0.1 | < 0.01 |  |  | 16
| 1 | 4 | < 0.01 | < 0.05 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | 0.02684 | < 0.05 | < 0.05 | < 0.5 | < 0.1 | < 0.01 |  |  | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.033446 | < 0.05 | < 0.05 | < 0.5 | < 0.1 | < 0.01 |  |  | 109
| 1 | 6 | < 0.01 | < 0.05 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | 0.023437 | < 0.05 | < 0.05 | < 0.5 | < 0.1 | < 0.01
| Mean |  | 1.07: 0.01; 1.11: 0.01; 1.12: 0.01; 1.14: 0.029032
| Median |  | 1.07: 0.01; 1.11: 0.01; 1.12: 0.01; 1.14: 0.027034
| Std Dev. |  | 1.07: 0; 1.11: 0; 1.12: 0; 1.14: 0.0058959
| Rel.Std.Dev. |  | 1.07: 0%; 1.11: 0%; 1.12: 0%; 1.14: 20.31%
PDM3
| Z-Score (Absolute) |  | 0.67 | 2.02 | 0 | 0.67 | 0 | 0.11 | 0.67 | 2.57 | 2.02 | 2.02 | 32.37 | 5.39 | 0.67
| Analytical results for Tb in OREAS 230b (Indicative Value 0.41 ppm) | [16 empty] | Indicative
Tb | Round | Replicate | 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.10: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07; 1.07: 1.19; 1.10: 1.22 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES/MS; 1.10: AR*MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g; 1.07: 0.25g; 1.10: 01g | 2
| 1 | 1 | 1.05: 0.33134; 1.06: 0.4; 1.07: 0.3; 1.10: 0.6 | 1
| 1 | 2 | 1.05: 0.32233; 1.06: 0.4; 1.07: 0.3; 1.10: 0.6 | 4
| 1 | 3 | 1.05: 0.32777; 1.06: 0.4; 1.07: 0.29; 1.10: 0.6 | 16
| 1 | 4 | 1.05: 0.34138; 1.06: 0.4; 1.07: 0.29; 1.10: 0.6 | 0.40605
| 1 | 5 | 1.05: 0.31808; 1.06: 0.4; 1.07: 0.29; 1.10: 0.6 | 10
| 1 | 6 | 1.05: 0.34423; 1.06: 0.4; 1.07: 0.29; 1.10: 0.6
| Mean |  | 1.05: 0.33086; 1.06: 0.4; 1.07: 0.29333; 1.10: 0.6
| Median |  | 1.05: 0.32956; 1.06: 0.4; 1.07: 0.29; 1.10: 0.6
| Std Dev. |  | 1.05: 0.010346; 1.06: 6.0809e-17; 1.07: 0.005164; 1.10: 0
| Rel.Std.Dev. |  | 1.05: 3.127%; 1.06: 1.52e-14%; 1.07: 1.76%; 1.10: 0%
| PDM3 |  | 1.05: -18.52%; 1.06: -1.489%; 1.07: -27.76%; 1.10: 47.77%
| Z-Score (Absolute) |  | 1.05: 0.44; 1.06: 0.44; 1.07: 0.91; 1.10: 2.97
| Analytical results for Te in OREAS 230b (Certified Value 0.1 ppm) | [16 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 0.254 | < 0.2 | 0.09 | 0.1 | 0.11 | 0.13 | 0.11 | 0.1 | 0.11149 | 0.09 | 0.11 | 0.08 | < 0.5 | 0.1 | 0.1 | 1
| 1 | 2 |  | < 0.2 | 0.08 | 0.11 | 0.11 | 0.08 | 0.11 | 0.1 | 0.06892 | 0.09 | 0.1 | 0.08 | < 0.5 | 0.1 | 0.12 | 23
| 1 | 3 | 0.071 | < 0.2 | 0.09 | 0.08 | 0.11 | 0.1 | 0.11 | 0.11 | 0.10377 | 0.07 | 0.12 | 0.05 | < 0.5 | 0.2 | 0.12 | 16
| 1 | 4 | 0.037 | < 0.2 | 0.09 | 0.1 | 0.11 | 0.08 | 0.11 | 0.11 | 0.098471 | 0.09 | 0.11 | 0.08 | < 0.5 | 0.1 | 0.09 | 0.10263
| 1 | 5 | 0.243 | < 0.2 | 0.11 | 0.1 | 0.1 | 0.11 | 0.12 | 0.09 | 0.11684 | 0.06 | 0.11 | 0.03 | < 0.5 | 0.1 | 0.13 | 110
| 1 | 6 |  | < 0.2 | 0.09 | 0.1 | 0.1 | 0.1 | 0.11 | 0.11 | 0.10445 | 0.09 | 0.1 | 0.12 | < 0.5 | 0.1 | 0.12
| Mean |  | 0.15125 |  | 0.091667 | 0.098333 | 0.10667 | 0.1 | 0.11167 | 0.10333 | 0.10066 | 0.081667 | 0.10833 | 0.073333 |  | 0.11667 | 0.11333
| Median |  | 0.157 |  | 0.09 | 0.1 | 0.11 | 0.1 | 0.11 | 0.105 | 0.10411 | 0.09 | 0.11 | 0.08 |  | 0.1 | 0.12
| Std Dev. |  | 0.11324 |  | 0.0098319 | 0.0098319 | 0.005164 | 0.018974 | 0.0040825 | 0.008165 | 0.016824 | 0.013292 | 0.0075277 | 0.030768 |  | 0.040825 | 0.015055
| Rel.Std.Dev. |  | 74.87% |  | 10.73% | 9.999% | 4.841% | 18.97% | 3.656% | 7.902% | 16.71% | 16.28% | 6.949% | 41.96% |  | 34.99% | 13.28%
| PDM3 |  | 47.37% |  | -10.69% | -4.19% | 3.93% | -2.566% | 8.801% | 0.6818% | -1.927% | -20.43% | 5.553% | -28.55% |  | 13.67% | 10.43%
| Z-Score (Absolute) |  | 4.15 | 0.17 | 0.87 | 0.31 | 0.39 | 0.17 | 0.81 | 0.11 | 0.11 | 1.71 | 0.53 | 2.41 | 12.47 | NA | 0.96
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 230b (Certified Value 0.35 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 0.364 | 0.36977 | 0.4 | 0.3 | 0.4 | 0.4 | 0.4 | 0.3 | 0.4 | 0.3 | 0.6 | 0.3 | 0.3 |  |  | 1
| 1 | 2 | 0.376 | 0.36184 | 0.4 | 0.3 | 0.4 | 0.4 | 0.4 | 0.3 | 0.4 | 0.3 | 0.6 | 0.3 | 0.3 |  |  | 24
| 1 | 3 | 0.37 | 0.36204 | 0.4 | 0.3 | 0.3 | 0.3 | 0.4 | 0.3 | 0.4 | 0.3 | 0.6 | 0.3 | 0.3 |  |  | 16
| 1 | 4 | 0.383 | 0.38136 | 0.4 | 0.3 | 0.4 | 0.4 | 0.4 | 0.4 | 0.41 | 0.3 | 0.6 | 0.3 | 0.4 |  |  | 0.35382
| 1 | 5 | 0.391 | 0.36546 | 0.4 | 0.3 | 0.3 | 0.4 | 0.4 | 0.4 | 0.4 | 0.3 | 0.6 | 0.3 | 0.4 |  |  | 111
| 1 | 6 | 0.369 | 0.37159 | 0.4 | 0.3 | 0.3 | 0.4 | 0.4 | 0.3 | 0.4 | 0.3 | 0.6 | 0.3 | 0.3
| Mean |  | 0.3755 | 0.36868 | 0.4 | 0.3 | 0.35 | 0.38333 | 0.4 | 0.33333 | 0.40167 | 0.3 | 0.6 | 0.3 | 0.33333
| Median |  | 0.373 | 0.36761 | 0.4 | 0.3 | 0.35 | 0.4 | 0.4 | 0.3 | 0.4 | 0.3 | 0.6 | 0.3 | 0.3
| Std Dev. |  | 0.010015 | 0.0073738 | 6.0809e-17 | 0 | 0.054772 | 0.040825 | 6.0809e-17 | 0.05164 | 0.0040825 | 0 | 0 | 0 | 0.05164
| Rel.Std.Dev. |  | 2.667% | 2% | 1.52e-14% | 0% | 15.65% | 10.65% | 1.52e-14% | 15.49% | 1.016% | 0% | 0% | 0% | 15.49%
| PDM3 |  | 6.127% | 4.199% | 13.05% | -15.21% | -1.08% | 8.341% | 13.05% | -5.79% | 13.52% | -15.21% | 69.58% | -15.21% | -5.79%
| Z-Score (Absolute) |  | 0.14 | 0 | 0.64 | 1.4 | 0.38 | 0.3 | 0.64 | 0.72 | 0.67 | 1.4 | 4.73 | 1.4 | 0.72
| Analytical results for Ti in OREAS 230b (Certified Value 0.359 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.18889 | 0.39507 | 0.37 | 0.352 | 0.35 | 0.369 | 0.43 | 0.417 | 0.36545 | > 0.5 | 0.31 | 0.46 | 0.407 | 0.26 |  | 1
| 1 | 2 | 0.20408 | 0.39479 | 0.35 | 0.357 | 0.353 | 0.377 | 0.433 | 0.395 | 0.36255 | > 0.5 | 0.31 | 0.46 | 0.416 | 0.24 |  | 25
| 1 | 3 | 0.21547 | 0.39896 | 0.36 | 0.357 | 0.361 | 0.369 | 0.438 | 0.417 | 0.36374 | > 0.5 | 0.31 | 0.45 | 0.411 | 0.25 |  | 16
| 1 | 4 | 0.20969 | 0.39791 | 0.35 | 0.338 | 0.337 | 0.393 | 0.435 | 0.397 | 0.32758 | > 0.5 | 0.31 | 0.46 | 0.423 | 0.24 |  | 0.35895
| 1 | 5 | 0.2228 | 0.38481 | 0.36 | 0.353 | 0.342 | 0.406 | 0.424 | 0.394 | 0.33136 | > 0.5 | 0.31 | 0.45 | 0.421 | 0.25 |  | 112
| 1 | 6 | 0.20311 | 0.39631 | 0.35 | 0.361 | 0.348 | 0.394 | 0.442 | 0.41 | 0.34778 | > 0.5 | 0.31 | 0.46 | 0.413 | 0.25
| Mean |  | 0.20734 | 0.39464 | 0.35667 | 0.353 | 0.3485 | 0.38467 | 0.43367 | 0.405 | 0.34974 |  | 0.31 | 0.45667 | 0.41517 | 0.24833
| Median |  | 0.20689 | 0.39569 | 0.355 | 0.355 | 0.349 | 0.385 | 0.434 | 0.4035 | 0.35516 |  | 0.31 | 0.46 | 0.4145 | 0.25
| Std Dev. |  | 0.011655 | 0.0050796 | 0.008165 | 0.008025 | 0.0084083 | 0.01524 | 0.0062823 | 0.010936 | 0.016968 |  | 0 | 0.005164 | 0.00608 | 0.0075277
| Rel.Std.Dev. |  | 5.621% | 1.287% | 2.289% | 2.273% | 2.413% | 3.962% | 1.449% | 2.7% | 4.851% |  | 0% | 1.131% | 1.464% | 3.031%
| PDM3 |  | -42.24% | 9.942% | -0.637% | -1.659% | -2.912% | 7.163% | 20.81% | 12.83% | -2.566% |  | -13.64% | 27.22% | 15.66% | -30.82%
| Z-Score (Absolute) |  | 2.16 | 0.55 | 0 | 0.05 | 0.12 | 0.4 | 1.11 | 0.7 | 0.1 | NA | 0.67 | 1.44 | 0.85 | 1.57
| Analytical results for Tl in OREAS 230b (Certified Value 0.045 ppm) | [16 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.026 | < 0.1 | 0.05 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.028181 | 0.05 | 0.04 | 0.051 | < 1 | 0.06 | 0.03 | 1
| 1 | 2 | 0.026 | < 0.1 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.020792 | 0.05 | 0.04 | 0.045 | < 1 | < 0.05 | 0.03 | 26
| 1 | 3 | 0.025 | < 0.1 | 0.04 | 0.05 | 0.04 | 0.04 | 0.05 | 0.04 | 0.023783 | 0.05 | 0.04 | 0.049 | < 1 | 0.06 | 0.03 | 16
| 1 | 4 | 0.026 | < 0.1 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.05 | 0.024369 | 0.05 | 0.04 | 0.047 | < 1 | 0.06 | 0.03 | 0.045241
| 1 | 5 | 0.023 | < 0.1 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.04 | 0.021417 | 0.05 | 0.04 | 0.052 | < 1 | < 0.05 | 0.03 | 113
| 1 | 6 | 0.026 | < 0.1 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.02045 | 0.05 | 0.04 | 0.049 | < 1 | 0.06 | 0.03
| Mean |  | 0.025333 |  | 0.045 | 0.043333 | 0.045 | 0.043333 | 0.046667 | 0.045 | 0.023165 | 0.05 | 0.04 | 0.048833 |  | 0.06 | 0.03
| Median |  | 0.026 |  | 0.045 | 0.04 | 0.045 | 0.04 | 0.05 | 0.045 | 0.0226 | 0.05 | 0.04 | 0.049 |  | 0.06 | 0.03
| Std Dev. |  | 0.0012111 |  | 0.0054772 | 0.005164 | 0.0054772 | 0.005164 | 0.005164 | 0.0054772 | 0.0029343 | 7.6012e-18 | 0 | 0.0025626 |  | 0 | 0
| Rel.Std.Dev. |  | 4.781% |  | 12.17% | 11.92% | 12.17% | 11.92% | 11.07% | 12.17% | 12.67% | 1.52e-14% | 0% | 5.248% |  | 0% | 0%
| PDM3 |  | -44% |  | -0.5321% | -4.216% | -0.5321% | -4.216% | 3.152% | -0.5321% | -48.8% | 10.52% | -11.58% | 7.941% |  | 32.62% | -33.69%
| Z-Score (Absolute) |  | 3.46 | 0.88 | 0 | 0.29 | 0 | 0.29 | 0.29 | 0 | 3.84 | 0.88 | 0.88 | 0.67 | 80.04 | 0.59 | 2.64
| Analytical results for Tm in OREAS 230b (Indicative Value 0.24 ppm) | [16 empty] | Indicative
Tm | Round | Replicate | 1.05: Lab; 1.06: Lab | 1
| No | No | 1.05: 1.06; 1.06: 1.07 | ppm
|  |  | 1.05: AR*MS; 1.06: AR*OES/MS | 2
|  |  | 1.05: 0.2g; 1.06: 0.5g | 2
| 1 | 1 | 1.05: 0.20999; 1.06: 0.3 | 1
| 1 | 2 | 1.05: 0.20826; 1.06: 0.3 | 5
| 1 | 3 | 1.05: 0.20702; 1.06: 0.2 | 16
| 1 | 4 | 1.05: 0.21147; 1.06: 0.3 | 0.23753
| 1 | 5 | 1.05: 0.20253; 1.06: 0.2 | 11
| 1 | 6 | 1.05: 0.21108; 1.06: 0.3
| Mean |  | 1.05: 0.20839; 1.06: 0.26667
| Median |  | 1.05: 0.20912; 1.06: 0.3
| Std Dev. |  | 1.05: 0.0033331; 1.06: 0.05164
| Rel.Std.Dev. |  | 1.05: 1.599%; 1.06: 19.36%
| PDM3 |  | 1.05: -12.27%; 1.06: 12.27%
| Z-Score (Absolute) |  | 1.05: 0.67; 1.06: 0.67
| Analytical results for U in OREAS 230b (Certified Value 0.11 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 0.088 | < 0.5 | 0.1 | 0.1 | 0.1 | 0.11 | 0.11 | 0.11 | 0.12 | 0.1 | 0.13 | < 0.5 | < 0.1 | 0.09 |  | 1
| 1 | 2 | 0.09 | < 0.5 | 0.1 | 0.1 | 0.1 | 0.1 | 0.12 | 0.11 | 0.12 | 0.1 | 0.12 | < 0.5 | < 0.1 | 0.09 |  | 28
| 1 | 3 | 0.088 | < 0.5 | 0.1 | 0.11 | 0.1 | 0.1 | 0.11 | 0.11 | 0.13 | 0.09 | 0.12 | < 0.5 | < 0.1 | 0.09 |  | 16
| 1 | 4 | 0.094 | < 0.5 | 0.1 | 0.1 | 0.1 | 0.12 | 0.12 | 0.12 | 0.12 | 0.09 | 0.12 | < 0.5 | < 0.1 | 0.09 |  | 0.10546
| 1 | 5 | 0.088 | < 0.5 | 0.1 | 0.1 | 0.1 | 0.12 | 0.12 | 0.11 | 0.12 | 0.1 | 0.12 | < 0.5 | < 0.1 | 0.09 |  | 114
| 1 | 6 | 0.088 | < 0.5 | 0.1 | 0.1 | 0.1 | 0.11 | 0.11 | 0.11 | 0.13 | 0.09 | 0.14 | < 0.5 | < 0.1 | 0.09
| Mean |  | 0.089333 |  | 0.1 | 0.10167 | 0.1 | 0.11 | 0.115 | 0.11167 | 0.12333 | 0.095 | 0.125 |  |  | 0.09
| Median |  | 0.088 |  | 0.1 | 0.1 | 0.1 | 0.11 | 0.115 | 0.11 | 0.12 | 0.095 | 0.12 |  |  | 0.09
| Std Dev. |  | 0.0024221 |  | 1.5202e-17 | 0.0040825 | 1.5202e-17 | 0.0089443 | 0.0054772 | 0.0040825 | 0.005164 | 0.0054772 | 0.0083666 |  |  | 1.5202e-17
| Rel.Std.Dev. |  | 2.711% |  | 1.52e-14% | 4.016% | 1.52e-14% | 8.131% | 4.763% | 3.656% | 4.187% | 5.766% | 6.693% |  |  | 1.689e-14%
| PDM3 |  | -15.29% |  | -5.178% | -3.597% | -5.178% | 4.304% | 9.045% | 5.885% | 16.95% | -9.919% | 18.53% |  |  | -14.66%
| Z-Score (Absolute) |  | 0.83 | 7.29 | 0.3 | 0.21 | 0.3 | 0.21 | 0.46 | 0.3 | 0.89 | 0.55 | 0.97 | 7.29 | 2.82 | 0.8
| Analytical results for V in OREAS 230b (Certified Value 169 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 0
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 0
| 1 | 1 | 113.146 | 168.2211 | 200 | 153 | 153 | 153 | 177 | 173 | 193 | 145 | 172.9 | 171 | 112 |  |  | 1
| 1 | 2 | 122.748 | 166.268 | 184 | 154 | 154 | 156 | 177 | 168 | 194 | 144 | 173.1 | 176 | 109 |  |  | 29
| 1 | 3 | 116.417 | 164.1032 | 188 | 154 | 156 | 157 | 179 | 174 | 196 | 145 | 169.3 | 174 | 109 |  |  | 16
| 1 | 4 | 117.868 | 165.8362 | 189 | 147 | 149 | 161 | 179 | 168 | 194 | 144 | 170.7 | 177 | 111 |  |  | 168.6942
| 1 | 5 | 123.581 | 163.2078 | 198 | 154 | 149 | 165 | 176 | 166 | 196 | 145 | 169.9 | 179 | 113 |  |  | 115
| 1 | 6 | 116.858 | 165.3805 | 194 | 153 | 152 | 163 | 181 | 170 | 195 | 143 | 170.3 | 173 | 112
| Mean |  | 118.4363 | 165.5028 | 192.1667 | 152.5 | 152.1667 | 159.1667 | 178.1667 | 169.8333 | 194.6667 | 144.3333 | 171.0333 | 175 | 111
| Median |  | 117.363 | 165.6084 | 191.5 | 153.5 | 152.5 | 159 | 178 | 169 | 194.5 | 144.5 | 170.5 | 175 | 111.5
| Std Dev. |  | 3.9984 | 1.7508 | 6.2102 | 2.7386 | 2.7869 | 4.5789 | 1.8348 | 3.1252 | 1.2111 | 0.8165 | 1.5933 | 2.8983 | 1.6733
| Rel.Std.Dev. |  | 3.376% | 1.058% | 3.232% | 1.796% | 1.831% | 2.877% | 1.03% | 1.84% | 0.6221% | 0.5657% | 0.9316% | 1.656% | 1.507%
| PDM3 |  | -29.79% | -1.892% | 13.91% | -9.6% | -9.797% | -5.648% | 5.615% | 0.6753% | 15.4% | -14.44% | 1.387% | 3.738% | -34.2%
| Z-Score (Absolute) |  | 2.44 | 0 | 1.38 | 0.67 | 0.69 | 0.33 | 0.66 | 0.22 | 1.51 | 1.1 | 0.29 | 0.49 | 2.83
| Analytical results for W in OREAS 230b (Certified Value 6.59 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  | 2
| 1 | 1 | 6.727 | 6.7334 | 5.9 | 6.31 | 6.21 | 7.13 | 7.52 | 6.47 | 2.1878 | 5.2 | 7.7 | 9.3 | 5.8 | 8.27 |  | 1
| 1 | 2 | 6.623 | 6.6189 | 5.3 | 6.48 | 6.27 | 6.47 | 7.39 | 6.44 | 1.8999 | 5.1 | 7.8 | 9 | 6.1 | 8.38 |  | 30
| 1 | 3 | 6.383 | 6.7192 | 5.4 | 6.45 | 6.23 | 6.25 | 7.64 | 6.51 | 2.226 | 5.2 | 7.6 | 9.2 | 6 | 8.48 |  | 16
| 1 | 4 | 6.192 | 6.634 | 5.4 | 6.26 | 6.12 | 6.98 | 7.07 | 6.34 | 1.7674 | 5.1 | 7.5 | 9.4 | 5.9 | 8.41 |  | 6.591
| 1 | 5 | 6.359 | 6.7004 | 5.6 | 6.41 | 6.32 | 7.22 | 7.49 | 6.37 | 2.099 | 5.3 | 7.7 | 9 | 6.1 | 8.43 |  | 116
| 1 | 6 | 6.327 | 6.6541 | 5.3 | 6.39 | 6.02 | 7.08 | 7.83 | 6.32 | 1.9777 | 5.2 | 7.3 | 9.5 | 6 | 8.42
| Mean |  | 6.4352 | 6.6767 | 5.4833 | 6.3833 | 6.195 | 6.855 | 7.49 | 6.4083 | 2.0263 | 5.1833 | 7.6 | 9.2333 | 5.9833 | 8.3983
| Median |  | 6.371 | 6.6772 | 5.4 | 6.4 | 6.22 | 7.03 | 7.505 | 6.405 | 2.0383 | 5.2 | 7.65 | 9.25 | 6 | 8.415
| Std Dev. |  | 0.1999 | 0.047451 | 0.23166 | 0.083825 | 0.10858 | 0.39733 | 0.25511 | 0.076267 | 0.17698 | 0.075277 | 0.17889 | 0.20656 | 0.1169 | 0.070828
| Rel.Std.Dev. |  | 3.106% | 0.7107% | 4.225% | 1.313% | 1.753% | 5.796% | 3.406% | 1.19% | 8.734% | 1.452% | 2.354% | 2.237% | 1.954% | 0.8434%
| PDM3 |  | -2.364% | 1.3% | -16.81% | -3.151% | -6.008% | 4.006% | 13.64% | -2.771% | -69.26% | -21.36% | 15.31% | 40.09% | -9.219% | 27.42%
| Z-Score (Absolute) |  | 0.01 | 0.25 | 0.92 | 0.04 | 0.22 | 0.42 | 1.05 | 0.01 | 4.31 | 1.21 | 1.15 | 2.75 | 0.43 | 1.94
| Analytical results for Y in OREAS 230b (Certified Value 12.1 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 1
| 1 | 1 | 9.395 | 12.6064 | 13.6 | 12.8 | 12 | 12.05 | 14.5 | 13.7 | 13.7 | 10.04 | 23.8 | 11.4 | 9.22 |  |  | 1
| 1 | 2 | 9.938 | 12.4898 | 12.6 | 13 | 12.25 | 12.3 | 14.4 | 13.5 | 14.1 | 10.09 | 23.7 | 11.6 | 9.04 |  |  | 31
| 1 | 3 | 10.284 | 12.5478 | 12.8 | 12.8 | 12.05 | 12 | 15.05 | 13.8 | 13.9 | 9.87 | 23.9 | 11.5 | 9.03 |  |  | 16
| 1 | 4 | 10.038 | 12.6176 | 13.1 | 12.8 | 12.2 | 12.5 | 14.3 | 13.45 | 13.9 | 9.78 | 24 | 11.5 | 9.09 |  |  | 12.1372
| 1 | 5 | 10.553 | 12.0704 | 13.6 | 12.85 | 12.1 | 12.6 | 14.45 | 13.05 | 14.2 | 9.85 | 23.5 | 11.5 | 9.24 |  |  | 117
| 1 | 6 | 9.999 | 12.3907 | 13.6 | 13.05 | 12.1 | 12.6 | 14.65 | 13.35 | 13.9 | 9.82 | 23.5 | 11.4 | 9.27
| Mean |  | 10.0345 | 12.4538 | 13.2167 | 12.8833 | 12.1167 | 12.3417 | 14.5583 | 13.475 | 13.95 | 9.9083 | 23.7333 | 11.4833 | 9.1483
| Median |  | 10.0185 | 12.5188 | 13.35 | 12.825 | 12.1 | 12.4 | 14.475 | 13.475 | 13.9 | 9.86 | 23.75 | 11.5 | 9.155
| Std Dev. |  | 0.38746 | 0.20554 | 0.44907 | 0.11255 | 0.093095 | 0.2691 | 0.26724 | 0.26599 | 0.17607 | 0.12608 | 0.20656 | 0.075277 | 0.10722
| Rel.Std.Dev. |  | 3.861% | 1.65% | 3.398% | 0.8736% | 0.7683% | 2.18% | 1.836% | 1.974% | 1.262% | 1.272% | 0.8703% | 0.6555% | 1.172%
| PDM3 |  | -17.32% | 2.608% | 8.894% | 6.147% | -0.1693% | 1.684% | 19.95% | 11.02% | 14.94% | -18.36% | 95.54% | -5.387% | -24.63%
| Z-Score (Absolute) |  | 1.6 | 0 | 0.5 | 0.28 | 0.22 | 0.07 | 1.39 | 0.67 | 0.99 | 1.68 | 7.45 | 0.64 | 2.18
| Analytical results for Yb in OREAS 230b (Indicative Value 1.2 ppm) | [16 empty] | Indicative
Yb | Round | Replicate | 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.10: Lab; 1.11: Lab | 1
| No | No | 1.05: 1.05; 1.06: 1.06; 1.07: 1.07; 1.10: 1.19; 1.11: 1.22 | ppm
|  |  | 1.05: AR*OES/MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.10: AR*OES/MS; 1.11: AR*MS | 2
|  |  | 1.05: 0.3g; 1.06: 0.2g; 1.07: 0.5g; 1.10: 0.25g; 1.11: 01g | 2
| 1 | 1 | 1.05: 0.926; 1.06: 1.3164; 1.07: 1.5; 1.10: 1.1; 1.11: < 0.5 | 1
| 1 | 2 | 1.05: 0.965; 1.06: 1.3568; 1.07: 1.4; 1.10: 1.1; 1.11: < 0.5 | 6
| 1 | 3 | 1.05: 0.934; 1.06: 1.3215; 1.07: 1.4; 1.10: 1.1; 1.11: < 0.5 | 16
| 1 | 4 | 1.05: 0.964; 1.06: 1.3194; 1.07: 1.5; 1.10: 1.1; 1.11: < 0.5 | 1.2003
| 1 | 5 | 1.05: 0.967; 1.06: 1.2375; 1.07: 1.5; 1.10: 1.1; 1.11: < 0.5 | 12
| 1 | 6 | 1.05: 0.945; 1.06: 1.254; 1.07: 1.5; 1.10: 1; 1.11: < 0.5
| Mean |  | 1.05: 0.95017; 1.06: 1.3009; 1.07: 1.4667; 1.10: 1.0833
| Median |  | 1.05: 0.9545; 1.06: 1.3179; 1.07: 1.5; 1.10: 1.1
| Std Dev. |  | 1.05: 0.017702; 1.06: 0.045478; 1.07: 0.05164; 1.10: 0.040825
| Rel.Std.Dev. |  | 1.05: 1.863%; 1.06: 3.496%; 1.07: 3.521%; 1.10: 3.768%
| PDM3 |  | 1.05: -20.84%; 1.06: 8.385%; 1.07: 22.19%; 1.10: -9.743%
| Z-Score (Absolute) |  | 1.05: 0.41; 1.06: 0.67; 1.07: 1.19; 1.10: 0; 1.11: 2.58
| Analytical results for Zn in OREAS 230b (Certified Value 96 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.05 | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 0.3g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.5g | 01g | 0.25g | 0.5g | 1
| 1 | 1 | 96.738 | 99.2103 | 98.5 | 98 | 96 | 101 | 99 | 92 | 103.5454 | 91 | 98 | 94 | 89 | 96 | 96.8 | 1
| 1 | 2 | 99.065 | 98.6119 | 93.7 | 98 | 98 | 101 | 97 | 91 | 93.0768 | 92 | 97 | 88 | 89.7 | 97 | 92.8 | 33
| 1 | 3 | 95.382 | 97.1792 | 94.3 | 105 | 97 | 95 | 97 | 93 | 101.3088 | 93 | 99 | 91 | 89.9 | 96 | 95.3 | 16
| 1 | 4 | 95.351 | 99.5747 | 95.4 | 96 | 96 | 103 | 98 | 94 | 100.3323 | 93 | 96 | 87 | 88.5 | 96 | 100 | 95.7961
| 1 | 5 | 95.927 | 97.3684 | 94.8 | 100 | 95 | 105 | 97 | 91 | 98.8208 | 94 | 98 | 96 | 86.9 | 96 | 95 | 118
| 1 | 6 | 92.999 | 97.9608 | 96.6 | 99 | 96 | 104 | 100 | 91 | 99.5993 | 93 | 100 | 86 | 89.1 | 96 | 94.1
| Mean |  | 95.9103 | 98.3175 | 95.55 | 99.3333 | 96.3333 | 101.5 | 98 | 92 | 99.4472 | 92.6667 | 98 | 90.3333 | 88.85 | 96.1667 | 95.6667
| Median |  | 95.6545 | 98.2863 | 95.1 | 98.5 | 96 | 102 | 97.5 | 91.5 | 99.9658 | 93 | 98 | 89.5 | 89.05 | 96 | 95.15
| Std Dev. |  | 1.9852 | 0.97857 | 1.7536 | 3.0768 | 1.0328 | 3.5637 | 1.2649 | 1.2649 | 3.5225 | 1.0328 | 1.4142 | 4.0332 | 1.0803 | 0.40825 | 2.5025
| Rel.Std.Dev. |  | 2.07% | 0.9953% | 1.835% | 3.097% | 1.072% | 3.511% | 1.291% | 1.375% | 3.542% | 1.115% | 1.443% | 4.465% | 1.216% | 0.4245% | 2.616%
| PDM3 |  | 0.1192% | 2.632% | -0.2569% | 3.692% | 0.5608% | 5.954% | 2.301% | -3.963% | 3.811% | -3.267% | 2.301% | -5.703% | -7.251% | 0.3868% | -0.1351%
| Z-Score (Absolute) |  | 0.08 | 0.67 | 0.19 | 0.99 | 0.05 | 1.67 | 0.57 | 1.31 | 1.03 | 1.1 | 0.57 | 1.83 | 2.29 | 0 | 0.16
| Analytical results for Zr in OREAS 230b (Certified Value 18.1 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.06 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 01g | 0.25g | 0.5g |  |  | 1
| 1 | 1 | 20.818 | 15 | 19.4 | 19 | 19.8 | 20.8 | 21.3 | 4.8461 | 29.5 | 14.3 | 76.8 | 18 | 13.2 |  |  | 1
| 1 | 2 | 20.6119 | 13.6 | 20.2 | 20 | 20.3 | 20.8 | 20.6 | 5.2946 | 28.1 | 14.7 | 76.8 | 18.2 | 12.8 |  |  | 34
| 1 | 3 | 20.9349 | 12.7 | 19.9 | 19.8 | 19.5 | 21.6 | 21.5 | 4.9972 | 29.4 | 14.6 | 76.4 | 18 | 12.9 |  |  | 16
| 1 | 4 | 20.3421 | 14.2 | 18.8 | 18.5 | 20.6 | 20.9 | 20.7 | 3.7108 | 29.7 | 14.8 | 77.3 | 17.3 | 12.9 |  |  | 18.1371
| 1 | 5 | 20.3512 | 15.5 | 19.8 | 19.1 | 20.8 | 21.3 | 21 | 5.6561 | 29.9 | 14.2 | 76.6 | 18 | 13.2 |  |  | 119
| 1 | 6 | 20.6674 | 12.6 | 20.4 | 19.1 | 20.8 | 21.4 | 21.2 | 5.11 | 28.6 | 14.1 | 78.2 | 17.6 | 13.2
| Mean |  | 20.6209 | 13.9333 | 19.75 | 19.25 | 20.3 | 21.1333 | 21.05 | 4.9358 | 29.2 | 14.45 | 77.0167 | 17.85 | 13.0333
| Median |  | 20.6396 | 13.9 | 19.85 | 19.1 | 20.45 | 21.1 | 21.1 | 5.0536 | 29.45 | 14.45 | 76.8 | 18 | 13.05
| Std Dev. |  | 0.2409 | 1.1894 | 0.57879 | 0.55408 | 0.54406 | 0.34448 | 0.35071 | 0.66185 | 0.69857 | 0.2881 | 0.65243 | 0.33317 | 0.18619
| Rel.Std.Dev. |  | 1.168% | 8.536% | 2.931% | 2.878% | 2.68% | 1.63% | 1.666% | 13.41% | 2.392% | 1.994% | 0.8471% | 1.866% | 1.429%
| PDM3 |  | 13.69% | -23.18% | 8.893% | 6.136% | 11.93% | 16.52% | 16.06% | -72.79% | 61% | -20.33% | 324.6% | -1.583% | -28.14%
| Z-Score (Absolute) |  | 0.31 | 2.06 | 0 | 0.18 | 0.2 | 0.49 | 0.46 | 5.26 | 3.35 | 1.88 | 20.32 | 0.67 | 2.38
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C987: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1059: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1077: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1095: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 230b (Indicative Value 13.11 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.1 | 1
| 1 | 2 | 13.12 | 10
| Mean |  | 13.11 | 16
| Median |  | 13.11 | 13.11
| Std Dev. |  | 0.014142 | 16
| Rel.Std.Dev. |  | 0.1079%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 230b (Indicative Value 10.38 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.36 | 1
| 1 | 2 | 10.4 | 11
| Mean |  | 10.38 | 16
| Median |  | 10.38 | 10.38
| Std Dev. |  | 0.028284 | 17
| Rel.Std.Dev. |  | 0.2725%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 230b (Indicative Value 13.26 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.24 | 1
| 1 | 2 | 13.27 | 12
| Mean |  | 13.255 | 16
| Median |  | 13.255 | 13.255
| Std Dev. |  | 0.021213 | 18
| Rel.Std.Dev. |  | 0.16%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 230b (Indicative Value 0.206 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.205 | 1
| 1 | 2 | 0.206 | 13
| Mean |  | 0.2055 | 16
| Median |  | 0.2055 | 0.2055
| Std Dev. |  | 0.00070711 | 19
| Rel.Std.Dev. |  | 0.3441%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 230b (Indicative Value 6.56 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.56 | 1
| 1 | 2 | 6.55 | 10
| Mean |  | 6.555 | 16
| Median |  | 6.555 | 6.555
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.1079%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 230b (Indicative Value 0.2 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 11
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 230b (Indicative Value 0.114 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.113 | 1
| 1 | 2 | 0.114 | 12
| Mean |  | 0.1135 | 16
| Median |  | 0.1135 | 0.1135
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.623%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 230b (Indicative Value 0.278 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.2783 | 1
| 1 | 2 | 0.2779 | 13
| Mean |  | 0.2781 | 16
| Median |  | 0.2781 | 0.27811
| Std Dev. |  | 0.00028284 | 19
| Rel.Std.Dev. |  | 0.1017%
| PDM3 |  | -0.004504%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 230b (Indicative Value 48.25 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 48.26 | 1
| 1 | 2 | 48.24 | 10
| Mean |  | 48.25 | 16
| Median |  | 48.25 | 48.25
| Std Dev. |  | 0.014142 | 16
| Rel.Std.Dev. |  | 0.02931%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 230b (Indicative Value 1.2 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 11
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 230b (Indicative Value 3.7 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.75 | 1
| 1 | 2 | 3.65 | 15
| Mean |  | 3.7 | 16
| Median |  | 3.7 | 3.7
| Std Dev. |  | 0.070711 | 21
| Rel.Std.Dev. |  | 1.911%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 230b (Indicative Value 0.375 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.38 | 17
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 230b (Indicative Value 0.305 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.31 | 13
| Mean |  | 0.305 | 16
| Median |  | 0.305 | 0.305
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 2.318%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 230b (Indicative Value 0.15 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.2 | 19
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 230b (Indicative Value 24.7 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 24.6 | 1
| 1 | 2 | 24.8 | 20
| Mean |  | 24.7 | 16
| Median |  | 24.7 | 24.7
| Std Dev. |  | 0.14142 | 26
| Rel.Std.Dev. |  | 0.5726%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 230b (Indicative Value 92 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 92.5 | 1
| 1 | 2 | 91 | 21
| Mean |  | 91.75 | 16
| Median |  | 91.75 | 91.75
| Std Dev. |  | 1.0607 | 27
| Rel.Std.Dev. |  | 1.156%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 230b (Indicative Value 0.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 22
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 230b (Indicative Value 0.22 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.22 | 23
| Mean |  | 0.22 | 16
| Median |  | 0.22 | 0.22
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 230b (Indicative Value 0.2 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 24
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 230b (Indicative Value 10.7 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.7 | 1
| 1 | 2 | 10.7 | 25
| Mean |  | 10.7 | 16
| Median |  | 10.7 | 10.7
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 230b (Indicative Value 47.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 47.2 | 1
| 1 | 2 | 47 | 26
| Mean |  | 47.1 | 16
| Median |  | 47.1 | 47.1
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 0.3003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 230b (Indicative Value 100 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 101 | 1
| 1 | 2 | 99 | 27
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 1.414%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 230b (Indicative Value 0.46 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.46 | 28
| Mean |  | 0.455 | 16
| Median |  | 0.455 | 0.455
| Std Dev. |  | 0.0070711 | 34
| Rel.Std.Dev. |  | 1.554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 230b (Indicative Value 186 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 182 | 1
| 1 | 2 | 190 | 29
| Mean |  | 186 | 16
| Median |  | 186 | 186
| Std Dev. |  | 5.6569 | 35
| Rel.Std.Dev. |  | 3.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 230b (Indicative Value 4.03 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.95 | 1
| 1 | 2 | 4.11 | 4
| Mean |  | 4.03 | 16
| Median |  | 4.03 | 4.03
| Std Dev. |  | 0.11314 | 36
| Rel.Std.Dev. |  | 2.807%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 230b (Indicative Value 2.67 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.66 | 1
| 1 | 2 | 2.68 | 5
| Mean |  | 2.67 | 16
| Median |  | 2.67 | 2.67
| Std Dev. |  | 0.014142 | 37
| Rel.Std.Dev. |  | 0.5297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 230b (Indicative Value 1 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 0.98 | 6
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0.028284 | 38
| Rel.Std.Dev. |  | 2.828%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 230b (Indicative Value 15.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.3 | 1
| 1 | 2 | 15.2 | 33
| Mean |  | 15.25 | 16
| Median |  | 15.25 | 15.25
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.4637%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 230b (Indicative Value 3.55 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.56 | 1
| 1 | 2 | 3.53 | 7
| Mean |  | 3.545 | 16
| Median |  | 3.545 | 3.545
| Std Dev. |  | 0.021213 | 40
| Rel.Std.Dev. |  | 0.5984%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 230b (Indicative Value 1.13 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.1 | 1
| 1 | 2 | 1.15 | 1
| Mean |  | 1.125 | 16
| Median |  | 1.125 | 1.125
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 230b (Indicative Value 1.94 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.94 | 1
| 1 | 2 | 1.93 | 19
| Mean |  | 1.935 | 16
| Median |  | 1.935 | 1.935
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.3654%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 230b (Indicative Value 0.93 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.94 | 1
| 1 | 2 | 0.92 | 3
| Mean |  | 0.93 | 16
| Median |  | 0.93 | 0.93
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 230b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 21
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 230b (Indicative Value 4.37 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.37 | 1
| 1 | 2 | 4.36 | 22
| Mean |  | 4.365 | 16
| Median |  | 4.365 | 4.365
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.162%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 230b (Indicative Value 0.38 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.37 | 4
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 230b (Indicative Value 0.162 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.16 | 1
| 1 | 2 | 0.163 | 24
| Mean |  | 0.1615 | 16
| Median |  | 0.1615 | 0.1615
| Std Dev. |  | 0.0021213 | 30
| Rel.Std.Dev. |  | 1.314%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 230b (Indicative Value 0.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 25
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 230b (Indicative Value 3.7 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.79 | 1
| 1 | 2 | 3.61 | 5
| Mean |  | 3.7 | 16
| Median |  | 3.7 | 3.7
| Std Dev. |  | 0.12728 | 32
| Rel.Std.Dev. |  | 3.44%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 230b (Indicative Value 8.25 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.29 | 1
| 1 | 2 | 8.21 | 6
| Mean |  | 8.25 | 16
| Median |  | 8.25 | 8.25
| Std Dev. |  | 0.056569 | 33
| Rel.Std.Dev. |  | 0.6857%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 230b (Indicative Value 74 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 74 | 1
| 1 | 2 | 74 | 28
| Mean |  | 74 | 16
| Median |  | 74 | 74
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 230b (Indicative Value 8 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8 | 1
| 1 | 2 | 8 | 29
| Mean |  | 8 | 16
| Median |  | 8 | 8
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 230b (Indicative Value 1.71 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.7 | 1
| 1 | 2 | 1.71 | 7
| Mean |  | 1.705 | 16
| Median |  | 1.705 | 1.705
| Std Dev. |  | 0.0070711 | 36
| Rel.Std.Dev. |  | 0.4147%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 230b (Indicative Value 4.6 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 4.6 | 31
| Mean |  | 4.6 | 16
| Median |  | 4.6 | 4.6
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 230b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 8
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 38
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 230b (Indicative Value 1.6 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.6 | 33
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 230b (Indicative Value 44.2 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 44.1 | 1
| 1 | 2 | 44.2 | 34
| Mean |  | 44.15 | 16
| Median |  | 44.15 | 44.15
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 0.1602%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 230b (Indicative Value 2.78 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.77 | 1
| 1 | 2 | 2.79 | 35
| Mean |  | 2.78 | 16
| Median |  | 2.78 | 2.78
| Std Dev. |  | 0.014142 | 41
| Rel.Std.Dev. |  | 0.5087%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 230b (Indicative Value 0.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.8 | 19
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 230b (Indicative Value 112 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 112 | 1
| 1 | 2 | 112 | 20
| Mean |  | 112 | 16
| Median |  | 112 | 112
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 230b (Indicative Value 0.65 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.68 | 1
| 1 | 2 | 0.62 | 21
| Mean |  | 0.65 | 16
| Median |  | 0.65 | 0.65
| Std Dev. |  | 0.042426 | 27
| Rel.Std.Dev. |  | 6.527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 230b (Indicative Value 0.63 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.63 | 1
| 1 | 2 | 0.63 | 4
| Mean |  | 0.63 | 16
| Median |  | 0.63 | 0.63
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 230b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 29
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 230b (Indicative Value 0.47 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.46 | 1
| 1 | 2 | 0.47 | 24
| Mean |  | 0.465 | 16
| Median |  | 0.465 | 0.465
| Std Dev. |  | 0.0070711 | 30
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 230b (Indicative Value 0.708 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.701 | 1
| 1 | 2 | 0.715 | 25
| Mean |  | 0.708 | 16
| Median |  | 0.708 | 0.708
| Std Dev. |  | 0.0098995 | 31
| Rel.Std.Dev. |  | 1.398%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 230b (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 26
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 32
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 230b (Indicative Value 0.38 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.37 | 5
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.0070711 | 33
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 230b (Indicative Value 0.13 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.13 | 1
| 1 | 2 | 0.13 | 28
| Mean |  | 0.13 | 16
| Median |  | 0.13 | 0.13
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 230b (Indicative Value 337 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 337 | 1
| 1 | 2 | 337 | 29
| Mean |  | 337 | 16
| Median |  | 337 | 337
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 230b (Indicative Value 9.23 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.15 | 1
| 1 | 2 | 9.3 | 30
| Mean |  | 9.225 | 16
| Median |  | 9.225 | 9.225
| Std Dev. |  | 0.10607 | 36
| Rel.Std.Dev. |  | 1.15%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 230b (Indicative Value 23.3 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23.1 | 1
| 1 | 2 | 23.4 | 31
| Mean |  | 23.25 | 16
| Median |  | 23.25 | 23.25
| Std Dev. |  | 0.21213 | 37
| Rel.Std.Dev. |  | 0.9124%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 230b (Indicative Value 2.68 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.77 | 1
| 1 | 2 | 2.59 | 6
| Mean |  | 2.68 | 16
| Median |  | 2.68 | 2.68
| Std Dev. |  | 0.12728 | 38
| Rel.Std.Dev. |  | 4.749%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 230b (Indicative Value 108 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 105 | 1
| 1 | 2 | 110 | 33
| Mean |  | 107.5 | 16
| Median |  | 107.5 | 107.5
| Std Dev. |  | 3.5355 | 39
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 230b (Indicative Value 67 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 66.5 | 1
| 1 | 2 | 66.5 | 34
| Mean |  | 66.5 | 16
| Median |  | 66.5 | 66.5
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.