## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 254c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 2.567 | 2.5324 | 2.6017 | 2.5583 | 2.5757
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 2.5456 | 2.5024 | 2.5888 | 2.5361 | 2.5551
Cyanide Leach
Au, Gold (ppm) | 2.4907 | 2.4107 | 2.5706 | 2.4874 | 2.4939
PhotonAssay (recommended gross mass* 470-500 g)
Au, Gold (ppm) | 2.5541 | 2.5171 | 2.5911 | 2.5516 | 2.5566
4-Acid Digestion
Ag, Silver (ppm) | 0.92721 | 0.83568 | 1.0187 | 0.88258 | 0.97184
Al, Aluminium (wt.%) | 5.0799 | 4.9543 | 5.2055 | 5.002 | 5.1578
As, Arsenic (ppm) | 17.5969 | 16.0793 | 19.1145 | 16.6487 | 18.5451
Ba, Barium (ppm) | 339.6477 | 328.7333 | 350.562 | 333.2453 | 346.05
Be, Beryllium (ppm) | 1.9334 | 1.8155 | 2.0512 | 1.8907 | 1.976
Bi, Bismuth (ppm) | 0.34686 | 0.3235 | 0.37022 | 0.321 | 0.37272
Ca, Calcium (wt.%) | 1.4637 | 1.4336 | 1.4937 | 1.4407 | 1.4867
Cd, Cadmium (ppm) | 0.071267 | 0.047587 | 0.094946 | 0.064054 | 0.07848
Ce, Cerium (ppm) | 69.5025 | 67.3274 | 71.6777 | 67.6937 | 71.3114
Co, Cobalt (ppm) | 20.097 | 19.2058 | 20.9882 | 19.6398 | 20.5542
Cr, Chromium (ppm) | 110.56 | 103.3536 | 117.7663 | 106.5175 | 114.6025
Cs, Caesium (ppm) | 2.4745 | 2.3704 | 2.5787 | 2.3757 | 2.5733
Cu, Copper (ppm) | 31.3785 | 29.5861 | 33.1708 | 30.4227 | 32.3342
Dy, Dysprosium (ppm) | 3.444 | 3.2374 | 3.6505 | 3.3531 | 3.5349
Er, Erbium (ppm) | 1.5928 | 1.4938 | 1.6918 | 1.5321 | 1.6535
Eu, Europium (ppm) | 1.5004 | 1.35 | 1.6509 | 1.4485 | 1.5524
Fe, Iron (wt.%) | 3.9332 | 3.8291 | 4.0373 | 3.865 | 4.0014
Ga, Gallium (ppm) | 15.037 | 14.1789 | 15.8951 | 14.5796 | 15.4945
Gd, Gadolinium (ppm) | 5.3145 | 4.7112 | 5.9177 | 5.097 | 5.532
Hf, Hafnium (ppm) | 4.4058 | 4.1338 | 4.6777 | 4.2682 | 4.5433
Ho, Holmium (ppm) | 0.61068 | 0.5526 | 0.66877 | 0.5802 | 0.64117
In, Indium (ppm) | 0.059506 | 0.046821 | 0.072191 | 0.055438 | 0.063574
K, Potassium (wt.%) | 1.2856 | 1.2493 | 1.3218 | 1.2626 | 1.3085
La, Lanthanum (ppm) | 37.8871 | 36.1845 | 39.5897 | 36.6817 | 39.0924
Li, Lithium (ppm) | 24.1261 | 23.0649 | 25.1873 | 23.6283 | 24.6239
Lu, Lutetium (ppm) | 0.17475 | 0.1483 | 0.20121 | IND | IND
Mg, Magnesium (wt.%) | 1.5093 | 1.4665 | 1.5521 | 1.4799 | 1.5387
Mn, Manganese (wt.%) | 0.045808 | 0.044734 | 0.046882 | 0.044975 | 0.04664
Mo, Molybdenum (ppm) | 3.8664 | 3.6043 | 4.1285 | 3.7461 | 3.9868
Na, Sodium (wt.%) | 0.9694 | 0.94116 | 0.99764 | 0.95229 | 0.98651
Nb, Niobium (ppm) | 30.5205 | 28.5724 | 32.4685 | 29.4598 | 31.5811
Nd, Neodymium (ppm) | 31.9647 | 26.8386 | 37.0909 | 30.7365 | 33.1929
Ni, Nickel (ppm) | 85.8479 | 82.7037 | 88.9921 | 84.2777 | 87.418
P, Phosphorus (wt.%) | 0.088594 | 0.086185 | 0.091003 | 0.0874 | 0.089788
Pb, Lead (ppm) | 10.7207 | 10.0193 | 11.4221 | 10.1759 | 11.2655
Pr, Praseodymium (ppm) | 8.491 | 7.8293 | 9.1528 | 8.2082 | 8.7738
Rb, Rubidium (ppm) | 64.5783 | 60.8772 | 68.2794 | 62.6817 | 66.4749
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.019788 | 0.018884 | 0.020692 | 0.018392 | 0.021184
Sb, Antimony (ppm) | 1.1652 | 1.0563 | 1.274 | 1.0708 | 1.2596
Sc, Scandium (ppm) | 10.1107 | 9.7489 | 10.4724 | 9.9114 | 10.3099
Sm, Samarium (ppm) | 6.0395 | 5.7065 | 6.3725 | 5.7551 | 6.3239
Sn, Tin (ppm) | 3.748 | 3.512 | 3.9839 | 3.5358 | 3.9602
Sr, Strontium (ppm) | 285.2934 | 273.6109 | 296.9759 | 279.8109 | 290.7759
Ta, Tantalum (ppm) | 1.9663 | 1.831 | 2.1016 | 1.8825 | 2.0502
Tb, Terbium (ppm) | 0.6762 | 0.61823 | 0.73416 | 0.6544 | 0.698
Th, Thorium (ppm) | 8.3684 | 7.87 | 8.8667 | 8.0945 | 8.6423
Ti, Titanium (wt.%) | 0.53161 | 0.51716 | 0.54606 | 0.5188 | 0.54442
Tl, Thallium (ppm) | 0.36087 | 0.34004 | 0.3817 | 0.34587 | 0.37587
Tm, Thulium (ppm) | 0.19532 | 0.17858 | 0.21206 | IND | IND
U, Uranium (ppm) | 1.5321 | 1.4183 | 1.6459 | 1.4756 | 1.5886
V, Vanadium (ppm) | 69.3031 | 67.1996 | 71.4067 | 67.4101 | 71.1961
W, Tungsten (ppm) | 3.203 | 3.0146 | 3.3914 | 3.0409 | 3.3651
Y, Yttrium (ppm) | 15.4241 | 14.7662 | 16.082 | 14.9347 | 15.9135
Yb, Ytterbium (ppm) | 1.2647 | 1.1578 | 1.3717 | 1.1862 | 1.3433
Zn, Zinc (ppm) | 74.5566 | 72.3362 | 76.7771 | 73.04 | 76.0733
Zr, Zirconium (ppm) | 188.7761 | 181.8567 | 195.6955 | 184.3627 | 193.1894
Aqua Regia Digestion
Ag, Silver (ppm) | 0.92813 | 0.89091 | 0.96536 | 0.90096 | 0.95531
Al, Aluminium (wt.%) | 1.6879 | 1.5842 | 1.7917 | 1.6352 | 1.7407
As, Arsenic (ppm) | 9.2498 | 8.5689 | 9.9307 | 8.9203 | 9.5792
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 94.5444 | 88.0062 | 101.0826 | 90.6434 | 98.4453
Be, Beryllium (ppm) | 1.0797 | 0.99187 | 1.1675 | 1.0401 | 1.1192
Bi, Bismuth (ppm) | 0.3104 | 0.28734 | 0.33346 | 0.29746 | 0.32334
Ca, Calcium (wt.%) | 0.4727 | 0.44003 | 0.50536 | 0.46241 | 0.48298
Cd, Cadmium (ppm) | 0.054611 | 0.037656 | 0.071566 | IND | IND
Ce, Cerium (ppm) | 39.7114 | 37.9071 | 41.5157 | 38.3545 | 41.0683
Co, Cobalt (ppm) | 16.7363 | 15.8469 | 17.6257 | 16.2401 | 17.2324
Cr, Chromium (ppm) | 60.4829 | 55.9983 | 64.9675 | 58.3243 | 62.6415
Cs, Caesium (ppm) | 0.92483 | 0.85013 | 0.99954 | 0.89596 | 0.95371
Cu, Copper (ppm) | 26.8533 | 25.0764 | 28.6302 | 25.9819 | 27.7246
Fe, Iron (wt.%) | 2.9783 | 2.8253 | 3.1314 | 2.913 | 3.0436
Ga, Gallium (ppm) | 6.0462 | 5.5703 | 6.5221 | 5.8623 | 6.2301
Ge, Germanium (ppm) | 0.10881 | 0.082905 | 0.13471 | IND | IND
Hg, Mercury (ppm) | 0.034979 | 0.01952 | 0.050438 | IND | IND
In, Indium (ppm) | 0.021595 | 0.018739 | 0.024452 | 0.018681 | 0.024509
K, Potassium (wt.%) | 0.29518 | 0.27254 | 0.31783 | 0.28592 | 0.30445
La, Lanthanum (ppm) | 22.6649 | 21.4427 | 23.887 | 21.9944 | 23.3353
Li, Lithium (ppm) | 8.728 | 8.132 | 9.324 | 8.4296 | 9.0264
Mg, Magnesium (wt.%) | 1.0266 | 0.96445 | 1.0888 | 1.0002 | 1.053
Mn, Manganese (wt.%) | 0.032684 | 0.031035 | 0.034334 | 0.03183 | 0.033539
Mo, Molybdenum (ppm) | 3.0641 | 2.8177 | 3.3105 | 2.9575 | 3.1708
Na, Sodium (wt.%) | 0.33497 | 0.31519 | 0.35475 | 0.3227 | 0.34724
Ni, Nickel (ppm) | 75.7506 | 71.8627 | 79.6384 | 73.5088 | 77.9924
P, Phosphorus (wt.%) | 0.051768 | 0.049377 | 0.05416 | 0.050696 | 0.052841
Pb, Lead (ppm) | 8.5537 | 7.991 | 9.1164 | 8.2228 | 8.8847
Rb, Rubidium (ppm) | 15.2838 | 14.3411 | 16.2265 | 14.7275 | 15.8401
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.019675 | 0.019118 | 0.020232 | 0.018738 | 0.020613
Sb, Antimony (ppm) | 0.5577 | 0.48468 | 0.63072 | 0.52976 | 0.58563
Sc, Scandium (ppm) | 3.9183 | 3.5833 | 4.2533 | 3.7582 | 4.0785
Sn, Tin (ppm) | 1.2599 | 1.1252 | 1.3947 | IND | IND
Sr, Strontium (ppm) | 59.5819 | 57.0161 | 62.1476 | 57.8383 | 61.3255
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.44055 | 0.38821 | 0.49289 | IND | IND
Th, Thorium (ppm) | 5.6198 | 5.2165 | 6.023 | 5.4281 | 5.8114
Ti, Titanium (wt.%) | 0.14718 | 0.13589 | 0.15848 | 0.14357 | 0.15079
Tl, Thallium (ppm) | 0.10789 | 0.096015 | 0.11976 | IND | IND
U, Uranium (ppm) | 0.91529 | 0.86048 | 0.97009 | 0.88616 | 0.94441
V, Vanadium (ppm) | 34.0941 | 32.2936 | 35.8947 | 32.9556 | 35.2327
W, Tungsten (ppm) | 0.56373 | 0.48591 | 0.64155 | 0.52602 | 0.60144
Y, Yttrium (ppm) | 8.9008 | 8.5324 | 9.2692 | 8.6612 | 9.1404
Yb, Ytterbium (ppm) | 0.64125 | 0.58622 | 0.69627 | IND | IND
Zn, Zinc (ppm) | 46.646 | 43.8888 | 49.4033 | 45.1742 | 48.1179
Zr, Zirconium (ppm) | 34.4295 | 26.1276 | 42.7314 | 31.6656 | 37.1934
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 254c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Ge | ppm | 0.13107 | Se | ppm | 0.99664
Hg | ppm | 0.058217 | Te | ppm | 0.071049
Aqua Regia Digestion
Dy | ppm | 2.162 | Lu | ppm | 0.09507 | Se | ppm | 0.27669
Er | ppm | 0.84583 | Nb | ppm | 0.68502 | Sm | ppm | 3.7828
Eu | ppm | 0.97834 | Nd | ppm | 21.5397 | Te | ppm | 0.040185
Gd | ppm | 3.2425 | Pd | ppb | < 10 | Tm | ppm | 0.10728
Hf | ppm | 0.66385 | Pr | ppm | 5.4797
Ho | ppm | 0.35808 | Pt | ppb | < 5
Borate Fusion XRF
Al2O3 | wt.% | 9.845 | MgO | wt.% | 2.57 | SiO2 | wt.% | 73.26
CaO | wt.% | 1.99 | MnO | wt.% | 0.059 | SO3 | wt.% | 0.054
Fe2O3 | wt.% | 5.73 | Na2O | wt.% | 1.31 | TiO2 | wt.% | 0.9305
K2O | wt.% | 1.53 | P2O5 | wt.% | 0.2005
Thermogravimetry
LOI1000 | wt.% | 2.315
Infrared Combustion
C | wt.% | 0.085 | S | wt.% | < 0.01
Laser Ablation ICP-MS
Ag | ppm | 1 | Hf | ppm | 6.095 | Sm | ppm | 6.17
As | ppm | 17.1 | Ho | ppm | 0.745 | Sn | ppm | 5
Ba | ppm | 342.5 | In | ppm | 0.0375 | Sr | ppm | 277.5
Be | ppm | 2.1 | La | ppm | 38.55 | Ta | ppm | 2.11
Bi | ppm | 0.35 | Lu | ppm | 0.22 | Tb | ppm | 0.75
Cd | ppm | 0.075 | Mn | wt.% | 0.047 | Te | ppm | < 0.2
Ce | ppm | 67.65 | Mo | ppm | 3.8 | Th | ppm | 8.755
Co | ppm | 20.45 | Nb | ppm | 32.2 | Ti | wt.% | 0.5625
Cr | ppm | 128 | Nd | ppm | 32.7 | Tl | ppm | 0.15
Cs | ppm | 2.45 | Ni | ppm | 90 | Tm | ppm | 0.265
Cu | ppm | 33 | Pb | ppm | 11 | U | ppm | 1.79
Dy | ppm | 4.01 | Pr | ppm | 8.855 | V | ppm | 72.9
Er | ppm | 1.97 | Rb | ppm | 62.3 | W | ppm | 3.75
Eu | ppm | 1.6 | Re | ppm | < 0.01 | Y | ppm | 19.1
Ga | ppm | 14.55 | Sb | ppm | 1.25 | Yb | ppm | 1.745
Gd | ppm | 5.08 | Sc | ppm | 9.35 | Zn | ppm | 75
Ge | ppm | 1.05 | Se | ppm | < 5 | Zr | ppm | 244
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 254c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 2.567 | 0.082378 | 2.4023 | 2.7318 | 2.3199 | 2.8141 | 3.209% | 6.418% | 9.627% | 2.4387 | 2.6954
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 2.5456 | 0.085869 | 2.3739 | 2.7173 | 2.288 | 2.8032 | 3.373% | 6.747% | 10.12% | 2.4183 | 2.6729
Cyanide Leach
Au, ppm | 2.4907 | 0.14502 | 2.2006 | 2.7807 | 2.0556 | 2.9257 | 5.822% | 11.64% | 17.47% | 2.3661 | 2.6152
PhotonAssay
Au, ppm | 2.5541 | 0.074616 | 2.4049 | 2.7033 | 2.3302 | 2.7779 | 2.921% | 5.843% | 8.764% | 2.4264 | 2.6818
4-Acid Digestion
Ag, ppm | 0.92721 | 0.070422 | 0.78637 | 1.0681 | 0.71595 | 1.1385 | 7.595% | 15.19% | 22.79% | 0.88085 | 0.97357
Al, wt.% | 5.0799 | 0.094492 | 4.8909 | 5.2689 | 4.7964 | 5.3634 | 1.86% | 3.72% | 5.58% | 4.8259 | 5.3339
As, ppm | 17.5969 | 1.3263 | 14.9442 | 20.2496 | 13.6179 | 21.5759 | 7.537% | 15.07% | 22.61% | 16.7171 | 18.4767
Ba, ppm | 339.6477 | 10.5418 | 318.5641 | 360.7312 | 308.0223 | 371.273 | 3.104% | 6.207% | 9.311% | 322.6653 | 356.63
Be, ppm | 1.9334 | 0.11714 | 1.6991 | 2.1676 | 1.5819 | 2.2848 | 6.059% | 12.12% | 18.18% | 1.8367 | 2.03
Bi, ppm | 0.34686 | 0.022696 | 0.30147 | 0.39225 | 0.27877 | 0.41495 | 6.543% | 13.09% | 19.63% | 0.32952 | 0.3642
Ca, wt.% | 1.4637 | 0.034129 | 1.3954 | 1.5319 | 1.3613 | 1.5661 | 2.332% | 4.663% | 6.995% | 1.3905 | 1.5369
Cd, ppm | 0.071267 | 0.014666 | 0.041935 | 0.1006 | 0.027269 | 0.11526 | 20.58% | 41.16% | 61.74% | 0.067703 | 0.07483
Ce, ppm | 69.5025 | 2.1876 | 65.1274 | 73.8777 | 62.9398 | 76.0653 | 3.147% | 6.295% | 9.442% | 66.0274 | 72.9777
Co, ppm | 20.097 | 1.1749 | 17.7472 | 22.4468 | 16.5722 | 23.6217 | 5.846% | 11.69% | 17.54% | 19.0921 | 21.1018
Cr, ppm | 110.56 | 7.3966 | 95.7668 | 125.3532 | 88.3702 | 132.7498 | 6.69% | 13.38% | 20.07% | 105.032 | 116.088
Cs, ppm | 2.4745 | 0.11737 | 2.2398 | 2.7093 | 2.1224 | 2.8266 | 4.743% | 9.486% | 14.23% | 2.3508 | 2.5982
Cu, ppm | 31.3785 | 2.6533 | 26.0718 | 36.6851 | 23.4184 | 39.3385 | 8.456% | 16.91% | 25.37% | 29.8095 | 32.9474
Dy, ppm | 3.444 | 0.10719 | 3.2296 | 3.6584 | 3.1224 | 3.7655 | 3.112% | 6.225% | 9.337% | 3.2718 | 3.6162
Er, ppm | 1.5928 | 0.055748 | 1.4813 | 1.7043 | 1.4255 | 1.76 | 3.5% | 7% | 10.5% | 1.5131 | 1.6724
Eu, ppm | 1.5004 | 0.11638 | 1.2677 | 1.7332 | 1.1513 | 1.8496 | 7.757% | 15.51% | 23.27% | 1.4254 | 1.5755
Fe, wt.% | 3.9332 | 0.085243 | 3.7627 | 4.1037 | 3.6775 | 4.1889 | 2.167% | 4.335% | 6.502% | 3.7366 | 4.1299
Ga, ppm | 15.037 | 1.2185 | 12.6 | 17.474 | 11.3815 | 18.6925 | 8.103% | 16.21% | 24.31% | 14.2852 | 15.7889
Gd, ppm | 5.3145 | 0.46496 | 4.3845 | 6.2444 | 3.9196 | 6.7093 | 8.749% | 17.5% | 26.25% | 5.0487 | 5.5802
Hf, ppm | 4.4058 | 0.3227 | 3.7603 | 5.0512 | 3.4376 | 5.3739 | 7.325% | 14.65% | 21.97% | 4.1855 | 4.626
Ho, ppm | 0.61068 | 0.046441 | 0.5178 | 0.70357 | 0.47136 | 0.75001 | 7.605% | 15.21% | 22.81% | 0.58015 | 0.64122
In, ppm | 0.059506 | 0.007306 | 0.044894 | 0.074118 | 0.037588 | 0.081424 | 12.28% | 24.56% | 36.83% | 0.056531 | 0.062481
K, wt.% | 1.2856 | 0.047133 | 1.1913 | 1.3798 | 1.1442 | 1.427 | 3.666% | 7.333% | 11% | 1.2213 | 1.3499
La, ppm | 37.8871 | 1.4889 | 34.9093 | 40.8648 | 33.4205 | 42.3536 | 3.93% | 7.859% | 11.79% | 35.9927 | 39.7814
Li, ppm | 24.1261 | 1.4324 | 21.2612 | 26.991 | 19.8288 | 28.4234 | 5.937% | 11.87% | 17.81% | 22.9198 | 25.3324
Lu, ppm | 0.17475 | 0.016806 | 0.14114 | 0.20837 | 0.12434 | 0.22517 | 9.617% | 19.23% | 28.85% | 0.16602 | 0.18349
Mg, wt.% | 1.5093 | 0.028167 | 1.453 | 1.5656 | 1.4248 | 1.5938 | 1.866% | 3.732% | 5.599% | 1.4338 | 1.5848
Mn, wt.% | 0.045808 | 0.0011136 | 0.043581 | 0.048035 | 0.042467 | 0.049149 | 2.431% | 4.862% | 7.293% | 0.043517 | 0.048098
Mo, ppm | 3.8664 | 0.22802 | 3.4104 | 4.3225 | 3.1824 | 4.5505 | 5.897% | 11.79% | 17.69% | 3.6731 | 4.0597
Na, wt.% | 0.9694 | 0.023324 | 0.92275 | 1.016 | 0.89943 | 1.0394 | 2.406% | 4.812% | 7.218% | 0.92093 | 1.0179
Nb, ppm | 30.5205 | 2.3045 | 25.9114 | 35.1295 | 23.6069 | 37.434 | 7.551% | 15.1% | 22.65% | 28.9944 | 32.0465
Nd, ppm | 31.9647 | 4.3922 | 23.1803 | 40.7491 | 18.7881 | 45.1414 | 13.74% | 27.48% | 41.22% | 30.3665 | 33.563
Ni, ppm | 85.8479 | 3.2552 | 79.3375 | 92.3583 | 76.0823 | 95.6135 | 3.792% | 7.584% | 11.38% | 81.5555 | 90.1403
P, wt.% | 0.088594 | 0.0023359 | 0.083922 | 0.093266 | 0.081586 | 0.095602 | 2.637% | 5.273% | 7.91% | 0.084164 | 0.093024
Pb, ppm | 10.7207 | 0.90291 | 8.9149 | 12.5265 | 8.012 | 13.4294 | 8.422% | 16.84% | 25.27% | 10.1847 | 11.2567
Pr, ppm | 8.491 | 0.40462 | 7.6818 | 9.3003 | 7.2772 | 9.7049 | 4.765% | 9.531% | 14.3% | 8.0665 | 8.9156
Rb, ppm | 64.5783 | 4.3889 | 55.8005 | 73.3561 | 51.4116 | 77.745 | 6.796% | 13.59% | 20.39% | 61.3494 | 67.8072
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.019788 | 0.00094695 | 0.017894 | 0.021682 | 0.016947 | 0.022629 | 4.786% | 9.571% | 14.36% | 0.018798 | 0.020777
Sb, ppm | 1.1652 | 0.071132 | 1.0229 | 1.3075 | 0.9518 | 1.3786 | 6.105% | 12.21% | 18.31% | 1.1069 | 1.2235
Sc, ppm | 10.1107 | 0.35187 | 9.4069 | 10.8144 | 9.0551 | 11.1663 | 3.48% | 6.96% | 10.44% | 9.6051 | 10.6162
Sm, ppm | 6.0395 | 0.29878 | 5.4419 | 6.6371 | 5.1431 | 6.9358 | 4.947% | 9.894% | 14.84% | 5.7375 | 6.3415
Sn, ppm | 3.748 | 0.17147 | 3.405 | 4.0909 | 3.2336 | 4.2624 | 4.575% | 9.15% | 13.73% | 3.5606 | 3.9354
Sr, ppm | 285.2934 | 12.5631 | 260.1671 | 310.4197 | 247.604 | 322.9828 | 4.404% | 8.807% | 13.21% | 271.0287 | 299.5581
Ta, ppm | 1.9663 | 0.14707 | 1.6722 | 2.2605 | 1.5251 | 2.4075 | 7.48% | 14.96% | 22.44% | 1.868 | 2.0646
Tb, ppm | 0.6762 | 0.040705 | 0.59479 | 0.75761 | 0.55408 | 0.79831 | 6.02% | 12.04% | 18.06% | 0.64239 | 0.71001
Th, ppm | 8.3684 | 0.5531 | 7.2622 | 9.4746 | 6.7091 | 10.0276 | 6.609% | 13.22% | 19.83% | 7.9499 | 8.7868
Ti, wt.% | 0.53161 | 0.014937 | 0.50174 | 0.56148 | 0.4868 | 0.57642 | 2.81% | 5.62% | 8.429% | 0.50503 | 0.55819
Tl, ppm | 0.36087 | 0.016919 | 0.32703 | 0.39471 | 0.31011 | 0.41163 | 4.688% | 9.377% | 14.07% | 0.34283 | 0.37892
Tm, ppm | 0.19532 | 0.01059 | 0.17414 | 0.2165 | 0.16355 | 0.22709 | 5.422% | 10.84% | 16.27% | 0.18555 | 0.20508
U, ppm | 1.5321 | 0.078672 | 1.3748 | 1.6895 | 1.2961 | 1.7681 | 5.135% | 10.27% | 15.4% | 1.4555 | 1.6087
V, ppm | 69.3031 | 1.8435 | 65.6161 | 72.9902 | 63.7725 | 74.8337 | 2.66% | 5.32% | 7.98% | 65.838 | 72.7683
W, ppm | 3.203 | 0.17719 | 2.8486 | 3.5574 | 2.6715 | 3.7346 | 5.532% | 11.06% | 16.6% | 3.0429 | 3.3632
Y, ppm | 15.4241 | 0.72136 | 13.9814 | 16.8668 | 13.26 | 17.5882 | 4.677% | 9.354% | 14.03% | 14.6529 | 16.1953
Yb, ppm | 1.2647 | 0.050881 | 1.163 | 1.3665 | 1.1121 | 1.4174 | 4.023% | 8.046% | 12.07% | 1.2015 | 1.328
Zn, ppm | 74.5566 | 2.6715 | 69.2137 | 79.8996 | 66.5422 | 82.5711 | 3.583% | 7.166% | 10.75% | 70.8288 | 78.2845
Zr, ppm | 188.7761 | 7.9147 | 172.9468 | 204.6054 | 165.0321 | 212.52 | 4.193% | 8.385% | 12.58% | 179.3373 | 198.2149
Aqua Regia Digestion
Ag, ppm | 0.92813 | 0.029507 | 0.86912 | 0.98715 | 0.83961 | 1.0167 | 3.179% | 6.358% | 9.538% | 0.88173 | 0.97454
Al, wt.% | 1.6879 | 0.15756 | 1.3728 | 2.0031 | 1.2153 | 2.1606 | 9.335% | 18.67% | 28% | 1.6036 | 1.7723
As, ppm | 9.2498 | 0.77175 | 7.7063 | 10.7933 | 6.9345 | 11.565 | 8.343% | 16.69% | 25.03% | 8.7873 | 9.7123
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 94.5444 | 6.5422 | 81.46 | 107.6287 | 74.9178 | 114.1709 | 6.92% | 13.84% | 20.76% | 89.8172 | 99.2716
Be, ppm | 1.0797 | 0.084299 | 0.91106 | 1.2483 | 0.82677 | 1.3326 | 7.808% | 15.62% | 23.42% | 1.0257 | 1.1336
Bi, ppm | 0.3104 | 0.024427 | 0.26155 | 0.35925 | 0.23712 | 0.38368 | 7.87% | 15.74% | 23.61% | 0.29488 | 0.32592
Ca, wt.% | 0.4727 | 0.051335 | 0.37003 | 0.57536 | 0.31869 | 0.6267 | 10.86% | 21.72% | 32.58% | 0.44906 | 0.49633
Cd, ppm | 0.054611 | 0.014548 | 0.025516 | 0.083706 | 0.010968 | 0.098254 | 26.64% | 53.28% | 79.92% | 0.05188 | 0.057341
Ce, ppm | 39.7114 | 1.8134 | 36.0847 | 43.3381 | 34.2713 | 45.1515 | 4.566% | 9.133% | 13.7% | 37.7258 | 41.697
Co, ppm | 16.7363 | 0.99006 | 14.7562 | 18.7164 | 13.7661 | 19.7065 | 5.916% | 11.83% | 17.75% | 15.8995 | 17.5731
Cr, ppm | 60.4829 | 7.4555 | 45.5718 | 75.394 | 38.1163 | 82.8495 | 12.33% | 24.65% | 36.98% | 57.4588 | 63.507
Cs, ppm | 0.92483 | 0.091938 | 0.74096 | 1.1087 | 0.64902 | 1.2007 | 9.941% | 19.88% | 29.82% | 0.87859 | 0.97108
Cu, ppm | 26.8533 | 2.2919 | 22.2694 | 31.4372 | 19.9775 | 33.7291 | 8.535% | 17.07% | 25.61% | 25.5106 | 28.196
Fe, wt.% | 2.9783 | 0.21445 | 2.5494 | 3.4072 | 2.335 | 3.6217 | 7.2% | 14.4% | 21.6% | 2.8294 | 3.1272
Ga, ppm | 6.0462 | 0.54475 | 4.9567 | 7.1357 | 4.4119 | 7.6804 | 9.01% | 18.02% | 27.03% | 5.7439 | 6.3485
Ge, ppm | 0.10881 | 0.023603 | 0.061604 | 0.15601 | 0.038002 | 0.17962 | 21.69% | 43.38% | 65.08% | 0.10337 | 0.11425
Hg, ppm | 0.034979 | 0.010059 | 0.014861 | 0.055098 | 0.0048016 | 0.065157 | 28.76% | 57.52% | 86.27% | 0.03323 | 0.036728
In, ppm | 0.021595 | 0.0020608 | 0.017474 | 0.025717 | 0.015413 | 0.027778 | 9.543% | 19.09% | 28.63% | 0.020515 | 0.022675
K, wt.% | 0.29518 | 0.035675 | 0.22383 | 0.36653 | 0.18816 | 0.40221 | 12.09% | 24.17% | 36.26% | 0.28042 | 0.30994
La, ppm | 22.6649 | 1.2532 | 20.1584 | 25.1713 | 18.9052 | 26.4246 | 5.529% | 11.06% | 16.59% | 21.5316 | 23.7981
Li, ppm | 8.728 | 0.6886 | 7.3508 | 10.1052 | 6.6622 | 10.7938 | 7.89% | 15.78% | 23.67% | 8.2916 | 9.1644
Mg, wt.% | 1.0266 | 0.077704 | 0.8712 | 1.182 | 0.7935 | 1.2597 | 7.569% | 15.14% | 22.71% | 0.97528 | 1.0779
Mn, wt.% | 0.032684 | 0.0020352 | 0.028614 | 0.036755 | 0.026579 | 0.03879 | 6.227% | 12.45% | 18.68% | 0.03105 | 0.034319
Mo, ppm | 3.0641 | 0.33425 | 2.3956 | 3.7326 | 2.0614 | 4.0669 | 10.91% | 21.82% | 32.73% | 2.9109 | 3.2173
Na, wt.% | 0.33497 | 0.024335 | 0.2863 | 0.38364 | 0.26196 | 0.40797 | 7.265% | 14.53% | 21.79% | 0.31822 | 0.35172
Ni, ppm | 75.7506 | 3.5331 | 68.6844 | 82.8168 | 65.1513 | 86.3499 | 4.664% | 9.328% | 13.99% | 71.9631 | 79.5381
P, wt.% | 0.051768 | 0.0030772 | 0.045614 | 0.057923 | 0.042537 | 0.061 | 5.944% | 11.89% | 17.83% | 0.04918 | 0.054357
Pb, ppm | 8.5537 | 0.65907 | 7.2356 | 9.8719 | 6.5765 | 10.5309 | 7.705% | 15.41% | 23.12% | 8.126 | 8.9814
Rb, ppm | 15.2838 | 0.86599 | 13.5518 | 17.0158 | 12.6858 | 17.8818 | 5.666% | 11.33% | 17% | 14.5196 | 16.048
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.019675 | 0.00065961 | 0.018356 | 0.020994 | 0.017696 | 0.021654 | 3.352% | 6.705% | 10.06% | 0.018691 | 0.020659
Sb, ppm | 0.5577 | 0.10235 | 0.35299 | 0.7624 | 0.25064 | 0.86476 | 18.35% | 36.71% | 55.06% | 0.52981 | 0.58558
Sc, ppm | 3.9183 | 0.37425 | 3.1698 | 4.6668 | 2.7956 | 5.0411 | 9.551% | 19.1% | 28.65% | 3.7224 | 4.1143
Sn, ppm | 1.2599 | 0.15344 | 0.95304 | 1.5668 | 0.7996 | 1.7202 | 12.18% | 24.36% | 36.54% | 1.1969 | 1.3229
Sr, ppm | 59.5819 | 2.8694 | 53.8431 | 65.3206 | 50.9738 | 68.19 | 4.816% | 9.632% | 14.45% | 56.6028 | 62.561
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.44055 | 0.037859 | 0.36483 | 0.51627 | 0.32697 | 0.55413 | 8.594% | 17.19% | 25.78% | 0.41852 | 0.46257
Th, ppm | 5.6198 | 0.43969 | 4.7404 | 6.4991 | 4.3007 | 6.9388 | 7.824% | 15.65% | 23.47% | 5.3388 | 5.9007
Ti, wt.% | 0.14718 | 0.015019 | 0.11715 | 0.17722 | 0.10213 | 0.19224 | 10.2% | 20.41% | 30.61% | 0.13983 | 0.15454
Tl, ppm | 0.10789 | 0.011428 | 0.085033 | 0.13074 | 0.073605 | 0.14217 | 10.59% | 21.18% | 31.78% | 0.10249 | 0.11328
U, ppm | 0.91529 | 0.066557 | 0.78217 | 1.0484 | 0.71561 | 1.115 | 7.272% | 14.54% | 21.82% | 0.86952 | 0.96105
V, ppm | 34.0941 | 2.173 | 29.7482 | 38.4401 | 27.5752 | 40.613 | 6.373% | 12.75% | 19.12% | 32.3894 | 35.7988
W, ppm | 0.56373 | 0.062648 | 0.43843 | 0.68903 | 0.37578 | 0.75167 | 11.11% | 22.23% | 33.34% | 0.53554 | 0.59192
Y, ppm | 8.9008 | 0.3406 | 8.2196 | 9.582 | 7.879 | 9.9226 | 3.827% | 7.653% | 11.48% | 8.4558 | 9.3459
Yb, ppm | 0.64125 | 0.041748 | 0.55775 | 0.72474 | 0.516 | 0.76649 | 6.51% | 13.02% | 19.53% | 0.60918 | 0.67331
Zn, ppm | 46.646 | 4.0758 | 38.4944 | 54.7977 | 34.4186 | 58.8735 | 8.738% | 17.48% | 26.21% | 44.3137 | 48.9784
Zr, ppm | 34.4295 | 13.5509 | 7.3277 | 61.5313 | 0 | 75.0823 | 39.36% | 78.72% | 118.1% | 32.708 | 36.151
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 254c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 254c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 254c (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 13 | 7 | 245440 | 0.084742 | 2.6702 | 2.7121
2 | 1 | ANSLu | 1 | 3 | 2 | 245441 | 0.086953 | 2.7261 | 2.7151
2 | 1 | ANSLu | 1 | 19 | 10 | 245442 | 0.085202 | 2.7043 | 2.7139
2 | 1 | ANSLu | 1 | 7 | 4 | 245443 | 0.088271 | 2.7066 | 2.714
2 | 1 | ANSLu | 1 | 17 | 9 | 245444 | 0.084313 | 2.6361 | 2.7103
2 | 1 | ANSLu | 1 | 5 | 3 | 245445 | 0.085952 | 2.6839 | 2.7128
2 | 1 | ANSLu | 1 | 20 | 10 | 245446 | 0.082741 | 2.7588 | 2.7168
2 | 1 | ANSLu | 1 | 8 | 4 | 245447 | 0.08497 | 2.7731 | 2.7176
2 | 1 | ANSLu | 1 | 16 | 8 | 245448 | 0.083459 | 2.6797 | 2.7126
2 | 1 | ANSLu | 1 | 10 | 5 | 245449 | 0.083921 | 2.7202 | 2.7147
2 | 1 | ANSLu | 1 | 15 | 8 | 245450 | 0.084222 | 2.6656 | 2.7118
2 | 1 | ANSLu | 1 | 9 | 5 | 245451 | 0.087349 | 2.621 | 2.7094
2 | 1 | ANSLu | 1 | 11 | 6 | 245452 | 0.087458 | 2.7352 | 2.7155
2 | 1 | ANSLu | 1 | 4 | 2 | 245453 | 0.08653 | 2.7782 | 2.7178
2 | 1 | ANSLu | 1 | 12 | 6 | 245454 | 0.087268 | 2.6508 | 2.711
2 | 1 | ANSLu | 1 | 18 | 9 | 245455 | 0.085875 | 2.6877 | 2.713
2 | 1 | ANSLu | 1 | 2 | 1 | 245456 | 0.085785 | 2.7529 | 2.7165
2 | 1 | ANSLu | 1 | 14 | 7 | 245457 | 0.085013 | 2.8342 | 2.7208
2 | 1 | ANSLu | 1 | 6 | 3 | 245458 | 0.083547 | 2.777 | 2.7178
2 | 1 | ANSLu | 1 | 1 | 1 | 245459 | 0.087883 | 2.7271 | 2.7151
Mean | [6 empty] | 0.085573 | 2.7144 | 2.7144
Median | [7 empty] | 2.7134 | 2.7144
SD | [7 empty] | 0.054415 | 0.0029062
RSD | [7 empty] | 2.005% | 0.1071%
RSD Upscale Factor | [7 empty] | 0.053408
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 254c (Certified Value 2.57 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*AAS | 2
|  |  | 0.085g | 25g | 50g | 30g | 50g | 30g | 50g | 50g | 50g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 2.7271 | 2.486 | 2.398 | 2.4637 | 2.62 | 2.62 | 2.54 | 2.58 | 2.626 | 2.43 | 2.56 | 2.71 | 2.52 | 2.712 | 2.546 | 2.6 | 2.5601 | 2.591 | 2.53 | 2.58 | 2.59 | 2.93 | 1
| 1 | 2 | 2.7529 | 2.633 | 2.373 | 2.4758 | 2.8 | 2.54 | 2.25 | 2.59 | 2.631 | 2.21 | 2.58 | 2.723 | 2.55 | 2.658 | 2.545 | 2.58 | 2.4845 | 2.603 | 2.54 | 2.64 | 2.58 | 2.95 | #N/A
| 1 | 3 | 2.7261 | 2.596 | 2.389 | 2.58 | 2.86 | 2.56 | 2.37 | 2.63 | 2.655 | 2.57 | 2.52 | 2.7018 | 2.53 | 2.718 | 2.545 | 2.6 | 2.5267 | 2.588 | 2.56 | 2.6 | 2.45 | 2.8 | 16
| 1 | 4 | 2.7782 | 2.605 | 2.433 | 2.5349 | 2.6 | 2.51 | 2.6 | 2.56 | 2.605 | 2.44 | 2.54 | 2.718 | 2.52 | 2.585 | 2.547 | 2.63 | 2.4828 | 2.587 | 2.57 | 2.6 | 2.56 | 2.76 | 2.567
| 1 | 5 | 2.6839 | 2.592 | 2.367 | 2.5363 | 2.74 | 2.57 | 2.64 | 2.59 | 2.639 | 2.43 | 2.58 | 2.698 | 2.54 | 2.692 | 2.545 | 2.64 | 2.4404 | 2.564 | 2.56 | 2.48 | 2.55 | 2.81 | 7
| 1 | 6 | 2.777 | 2.486 | 2.391 | 2.5004 | 2.59 | 2.58 | 2.63 | 2.56 | 2.647 | 2.45 | 2.52 | 2.697 | 2.53 | 2.703 | 2.547 | 2.63 | 2.534 | 2.596 | 2.54 | 2.57 | 2.55 | 2.92
|  | 7 | 2.7066
|  | 8 | 2.7731
|  | 9 | 2.621
|  | 10 | 2.7202
|  | 11 | 2.7352
|  | 12 | 2.6508
|  | 13 | 2.6702
|  | 14 | 2.8342
|  | 15 | 2.6656
|  | 16 | 2.6797
|  | 17 | 2.6361
|  | 18 | 2.6877
|  | 19 | 2.7043
|  | 20 | 2.7588
| Mean |  | 2.7144 | 2.5663 | 2.3918 | 2.5152 | 2.7017 | 2.5633 | 2.505 | 2.585 | 2.6338 | 2.4217 | 2.55 | 2.708 | 2.5317 | 2.678 | 2.5458 | 2.6133 | 2.5047 | 2.5882 | 2.55 | 2.5783 | 2.5467 | 2.8617
| Median |  | 2.7134 | 2.594 | 2.39 | 2.5177 | 2.68 | 2.565 | 2.57 | 2.585 | 2.635 | 2.435 | 2.55 | 2.7059 | 2.53 | 2.6975 | 2.5455 | 2.615 | 2.5056 | 2.5895 | 2.55 | 2.59 | 2.555 | 2.865
| Std Dev. |  | 0.054415 | 0.063852 | 0.023293 | 0.043497 | 0.11462 | 0.037238 | 0.15959 | 0.025884 | 0.0176 | 0.11669 | 0.027568 | 0.010845 | 0.01169 | 0.050251 | 0.00098319 | 0.023381 | 0.043449 | 0.013227 | 0.015492 | 0.053821 | 0.050067 | 0.08085
| Rel.Std.Dev. |  | 2.005% | 2.488% | 0.9739% | 1.729% | 4.242% | 1.453% | 6.371% | 1.001% | 0.6682% | 4.819% | 1.081% | 0.4005% | 0.4618% | 1.876% | 0.03862% | 0.8947% | 1.735% | 0.5111% | 0.6075% | 2.087% | 1.966% | 2.825%
| PDM3 |  | 5.743% | -0.0265% | -6.824% | -2.018% | 5.246% | -0.1434% | -2.416% | 0.7007% | 2.603% | -5.662% | -0.6628% | 5.491% | -1.377% | 4.324% | -0.8251% | 1.804% | -2.426% | 0.824% | -0.6628% | 0.441% | -0.7926% | 11.48%
| Z-Score (Absolute) |  | NA | 0.04 | 2.4 | 0.67 | 1.94 | 0 | 0.82 | 0.3 | 0.99 | 1.99 | 0.19 | 2.03 | 0.44 | 1.61 | 0.25 | 0.7 | 0.82 | 0.35 | 0.19 | 0.21 | 0.23 | 4.18
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 254c (Certified Value 2.55 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | 2
|  |  | 0.085g | 25g | 30g | 50g | 30g | 50g | 30g | 30g | 25g | 15g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 3
| 1 | 1 | 2.7271 | 2.59 | 2.4464 | 2.58 | 2.56 | 2.32 | 2.31 | 2.5972 | 2.61 | 2.5647 | 2.606 | 2.49 | 2.3396 | 2.5 | 2.53 | 2.53 | 2.504 | 1
| 1 | 2 | 2.7529 | 2.62 | 2.4211 | 2.62 | 2.54 | 2.34 | 2.23 | 2.603 | 2.54 | 2.4411 | 2.604 | 2.5 | 2.2646 | 2.65 | 2.68 | 2.41 | 2.369 | #N/A
| 1 | 3 | 2.7261 | 2.51 | 2.4373 | 2.8 | 2.55 | 2.4833 | 2.24 | 2.624 | 2.65 | 2.5853 | 2.603 | 2.51 | 2.3344 | 2.54 | 2.62 | 2.51 | 2.253 | 16
| 1 | 4 | 2.7782 | 2.58 | 2.4819 | 2.54 | 2.5 | 2.42 | 2.15 | 2.595 | 2.64 | 2.5338 | 2.606 | 2.51 | 2.3077 | 2.54 | 2.59 | 2.56 | 2.408 | 2.5456
| 1 | 5 | 2.6839 | 2.52 | 2.415 | 2.7 | 2.55 | 2.36 | 2.37 | 2.598 | 2.62 | 2.6471 | 2.604 | 2.52 | 2.2584 | 2.59 | 2.72 | 2.53 | 2.311 | 9
| 1 | 6 | 2.777 | 2.58 | 2.4393 | 2.54 | 2.51 | 2.41 | 2.16 | 2.5897 | 2.51 | 2.6368 | 2.603 | 2.49 | 2.3318 | 2.55 | 2.45 | 2.46 | 2.234
|  | 7 | 2.7066
|  | 8 | 2.7731
|  | 9 | 2.621
|  | 10 | 2.7202
|  | 11 | 2.7352
|  | 12 | 2.6508
|  | 13 | 2.6702
|  | 14 | 2.8342
|  | 15 | 2.6656
|  | 16 | 2.6797
|  | 17 | 2.6361
|  | 18 | 2.6877
|  | 19 | 2.7043
|  | 20 | 2.7588
| Mean |  | 2.7144 | 2.5667 | 2.4402 | 2.63 | 2.535 | 2.3889 | 2.2433 | 2.6011 | 2.595 | 2.5681 | 2.6043 | 2.5033 | 2.3061 | 2.5617 | 2.5983 | 2.5 | 2.3465
| Median |  | 2.7134 | 2.58 | 2.4383 | 2.6 | 2.545 | 2.385 | 2.235 | 2.5976 | 2.615 | 2.575 | 2.604 | 2.505 | 2.3197 | 2.545 | 2.605 | 2.52 | 2.34
| Std Dev. |  | 0.054415 | 0.04274 | 0.02359 | 0.10257 | 0.02429 | 0.060504 | 0.085245 | 0.012006 | 0.056833 | 0.075596 | 0.0013663 | 0.012111 | 0.036281 | 0.051929 | 0.098675 | 0.055136 | 0.10174
| Rel.Std.Dev. |  | 2.005% | 1.665% | 0.9667% | 3.9% | 0.9582% | 2.533% | 3.8% | 0.4616% | 2.19% | 2.944% | 0.05246% | 0.4838% | 1.573% | 2.027% | 3.798% | 2.205% | 4.336%
| PDM3 |  | 6.633% | 0.828% | -4.141% | 3.316% | -0.416% | -6.156% | -11.87% | 2.182% | 1.941% | 0.8856% | 2.308% | -1.66% | -9.409% | 0.6316% | 2.072% | -1.791% | -7.821%
| Z-Score (Absolute) |  | NA | 0.24 | 1.42 | 1.07 | 0.17 | 2.09 | 4 | 0.69 | 0.61 | 0.26 | 0.73 | 0.59 | 3.18 | 0.17 | 0.66 | 0.63 | 2.65
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 254c (Certified Value 2.49 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 05g | 50g | 50g | 200g | 50g | 30g | 30g | 30g | 30g | 20g | 05g | 3
| 1 | 1 | 2.7271 | 2.59 | 2.357 | 2.3318 | 2.53 | 2.57 | 2.3 | 2.4224 | 2.4 | 2.5 | 2.601 | 2.064 | 2.56 | 2.764 | 2.68 | 2.3005 | 2.573 | 1
| 1 | 2 | 2.7529 | 2.61 | 2.3326 | 2.3032 | 2.55 | 2.51 | 2.3 | 2.43 | 2.29 | 2.48 | 2.577 | 2.073 | 2.54 | 2.755 | 2.68 | 2.2508 | 2.533 | #N/A
| 1 | 3 | 2.7261 | 2.59 | 2.3482 | 2.3469 | 2.74 | 2.58 | 2.3 | 2.3843 | 2.23 | 2.49 | 2.595 | 2.071 | 2.56 | 2.753 | 2.67 | 2.232 | 2.563 | 16
| 1 | 4 | 2.7782 | 2.59 | 2.3911 | 2.2954 | 2.5 | 2.59 | 2.3 | 2.48 | 2.39 | 2.48 | 2.556 | 2.063 | 2.57 | 2.754 | 2.7 | 2.2917 | 2.587 | 2.4907
| 1 | 5 | 2.6839 | 2.61 | 2.3267 | 2.3376 | 2.62 | 2.49 | 2.4 | 2.51 | 2.37 | 2.49 | 2.568 | 2.089 | 2.53 | 2.763 | 2.58 | 2.251 | 2.582 | 11
| 1 | 6 | 2.777 | 2.57 | 2.3501 | 2.3765 | 2.49 | 2.51 | 2.3 | 2.5032 | 2.32 | 2.49 | 2.566 | 2.041 | 2.55 | 2.754 | 2.67 | 2.2384 | 2.565
|  | 7 | 2.7066
|  | 8 | 2.7731
|  | 9 | 2.621
|  | 10 | 2.7202
|  | 11 | 2.7352
|  | 12 | 2.6508
|  | 13 | 2.6702
|  | 14 | 2.8342
|  | 15 | 2.6656
|  | 16 | 2.6797
|  | 17 | 2.6361
|  | 18 | 2.6877
|  | 19 | 2.7043
|  | 20 | 2.7588
| Mean |  | 2.7144 | 2.5933 | 2.3509 | 2.3319 | 2.5717 | 2.5417 | 2.3167 | 2.455 | 2.3333 | 2.4883 | 2.5772 | 2.0668 | 2.5517 | 2.7572 | 2.6633 | 2.2607 | 2.5672
| Median |  | 2.7134 | 2.59 | 2.3492 | 2.3347 | 2.54 | 2.54 | 2.3 | 2.455 | 2.345 | 2.49 | 2.5725 | 2.0675 | 2.555 | 2.7545 | 2.675 | 2.2509 | 2.569
| Std Dev. |  | 0.054415 | 0.015055 | 0.022727 | 0.029684 | 0.094534 | 0.043089 | 0.040825 | 0.05032 | 0.065929 | 0.0075277 | 0.017566 | 0.015728 | 0.01472 | 0.0049565 | 0.042269 | 0.028481 | 0.019167
| Rel.Std.Dev. |  | 2.005% | 0.5805% | 0.9667% | 1.273% | 3.676% | 1.695% | 1.762% | 2.05% | 2.826% | 0.3025% | 0.6816% | 0.761% | 0.5769% | 0.1798% | 1.587% | 1.26% | 0.7466%
| PDM3 |  | 8.984% | 4.122% | -5.61% | -6.375% | 3.252% | 2.048% | -6.986% | -1.433% | -6.317% | -0.0938% | 3.473% | -17.02% | 2.449% | 10.7% | 6.932% | -9.232% | 3.071%
| Z-Score (Absolute) |  | NA | 0.47 | 0.98 | 1.09 | 0.34 | 0.16 | 1.18 | 0.36 | 1.08 | 0.16 | 0.37 | 2.67 | 0.22 | 1.44 | 0.88 | 1.51 | 0.31
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 254c (Certified Value 2.55 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | [11 empty] | 3
| 1 | 1 | 2.7271 | 2.612 | 2.65 | 2.57 | 2.52 | 2.59 | 2.53 | 2.54 | 2.496 | 2.609 | 2.486 | 2.532 | 1
| 1 | 2 | 2.7529 | 2.658 | 2.587 | 2.69 | 2.42 | 2.61 | 2.65 | 2.41 | 2.539 | 2.601 | 2.464 | 2.461 | #N/A
| 1 | 3 | 2.7261 | 2.662 | 2.59 | 2.54 | 2.58 | 2.65 | 2.55 | 2.53 | 2.616 | 2.527 | 2.405 | 2.532 | 16
| 1 | 4 | 2.7782 | 2.686 | 2.659 | 2.53 | 2.39 | 2.51 | 2.62 | 2.49 | 2.471 | 2.568 | 2.53 | 2.512 | 2.5541
| 1 | 5 | 2.6839 | 2.608 | 2.639 | 2.67 | 2.51 | 2.41 | 2.57 | 2.54 | 2.447 | 2.588 | 2.53 | 2.57 | 13
| 1 | 6 | 2.777 | 2.621 | 2.619 | 2.56 | 2.48 | 2.56 | 2.58 | 2.61 | 2.585 | 2.602 | 2.491 | 2.407
|  | 7 | 2.7066
|  | 8 | 2.7731
|  | 9 | 2.621
|  | 10 | 2.7202
|  | 11 | 2.7352
|  | 12 | 2.6508
|  | 13 | 2.6702
|  | 14 | 2.8342
|  | 15 | 2.6656
|  | 16 | 2.6797
|  | 17 | 2.6361
|  | 18 | 2.6877
|  | 19 | 2.7043
|  | 20 | 2.7588
| Mean |  | 2.7144 | 2.6412 | 2.624 | 2.5933 | 2.4833 | 2.555 | 2.5833 | 2.52 | 2.5257 | 2.5825 | 2.4843 | 2.5023
| Median |  | 2.7134 | 2.6395 | 2.629 | 2.565 | 2.495 | 2.575 | 2.575 | 2.535 | 2.5175 | 2.5945 | 2.4885 | 2.522
| Std Dev. |  | 0.054415 | 0.03189 | 0.030581 | 0.068896 | 0.069474 | 0.085264 | 0.044572 | 0.066332 | 0.066205 | 0.030807 | 0.046745 | 0.058702
| Rel.Std.Dev. |  | 2.005% | 1.207% | 1.165% | 2.657% | 2.798% | 3.337% | 1.725% | 2.632% | 2.621% | 1.193% | 1.882% | 2.346%
| PDM3 |  | 6.278% | 3.409% | 2.737% | 1.536% | -2.77% | 0.03559% | 1.145% | -1.335% | -1.113% | 1.112% | -2.731% | -2.026%
| Z-Score (Absolute) |  | NA | 1.52 | 1.21 | 0.67 | 1.26 | 0 | 0.5 | 0.62 | 0.52 | 0.48 | 1.24 | 0.93
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 254c (Certified Value 0.927 ppm) | [19 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.99 | 0.74971 | 0.7247 | 1 | 0.9 | 0.89 | 1.192 | 1 | 1.1 | 1.0329 | 0.94 | 0.77568 | 1.2 | 0.92 | 0.9 | 0.91 |  |  | 1
| 1 | 2 | 0.97 | 0.82968 | 0.6451 | 1 | 0.9 | 0.93 | 1.111 | 1.01 | 1 | 1.012 | 0.9 | 0.77261 | 1.2 | 0.92 | 0.9 | 1 |  |  | 20
| 1 | 3 | 0.94 | 0.91831 | 0.4822 | 1 | 0.9 | 0.91 | 1.1655 | 0.97 | 1.1 | 1.0184 | 0.94 | 0.76834 | 1.2 | 0.93 | 0.91 | 0.92 |  |  | 16
| 1 | 4 | 0.97 | 0.94213 | 0.5561 | 1 | 0.9 | 0.9 | 1.1338 | 1.04 | 1 | 1.0027 | 0.93 | 0.8348 | 1.3 | 0.89 | 0.9 | 1 |  |  | 0.92721
| 1 | 5 | 1 | 0.87832 | 0.5997 | 1 | 0.9 | 0.91 | 1.2248 | 0.92 | 0.9 | 0.99373 | 1.03 | 0.77651 | 1.2 | 0.9 | 0.89 | 0.97 |  |  | 15
| 1 | 6 | 0.92 | 0.95191 | 0.5074 | 0.9 | 0.9 | 0.89 | 1.1254 | 1.01 | 1 | 1.0293 | 0.98 | 0.83569 | 1.2 | 0.94 | 0.93 | 0.92
| Mean |  | 0.965 | 0.87834 | 0.58587 | 0.98333 | 0.9 | 0.905 | 1.1587 | 0.99167 | 1.0167 | 1.0148 | 0.95333 | 0.79394 | 1.2167 | 0.91667 | 0.905 | 0.95333
| Median |  | 0.97 | 0.89832 | 0.5779 | 1 | 0.9 | 0.905 | 1.1496 | 1.005 | 1 | 1.0152 | 0.94 | 0.7761 | 1.2 | 0.92 | 0.9 | 0.945
| Std Dev. |  | 0.030166 | 0.077452 | 0.090293 | 0.040825 | 0 | 0.015166 | 0.04367 | 0.041673 | 0.075277 | 0.015135 | 0.045461 | 0.032123 | 0.040825 | 0.018619 | 0.013784 | 0.041793
| Rel.Std.Dev. |  | 3.126% | 8.818% | 15.41% | 4.152% | 0% | 1.676% | 3.769% | 4.202% | 7.404% | 1.491% | 4.769% | 4.046% | 3.355% | 2.031% | 1.523% | 4.384%
| PDM3 |  | 4.076% | -5.27% | -36.81% | 6.053% | -2.935% | -2.395% | 24.97% | 6.952% | 9.648% | 9.449% | 2.817% | -14.37% | 31.22% | -1.137% | -2.395% | 2.817%
| Z-Score (Absolute) |  | 0.47 | 0.88 | 5.43 | NA | NA | 0.47 | 3.48 | 0.88 | NA | 1.24 | 0.29 | 2.2 | NA | 0.29 | 0.47 | 0.29
Results from laboratories 1.07, 1.09, 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 254c (Certified Value 5.08 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.2206 | 5.1175 | 5.3403 | 4.7547 | 5.06 | 5.41 | 5.14 | 5.06 | 5.25 | 5.18 | 4.8601 | 4.94 | 5.0966 | 5.2 | 5.07 | 5.1 | 4.93 | 5.09 | 1
| 1 | 2 | 5.2221 | 5.1291 | 5.2933 | 4.7384 | 5.12 | 5.43 | 5.18 | 4.95 | 5.11 | 5.18 | 4.9234 | 5.07 | 5.073 | 5.1 | 5.12 | 5.05 | 4.95 | 4.96 | #N/A
| 1 | 3 | 5.1591 | 5.1947 | 5.2767 | 4.7716 | 5.08 | 5.46 | 5.16 | 5.18 | 5.05 | 5.25 | 4.8417 | 5.02 | 5.0922 | 5.17 | 4.94 | 5.04 | 5 | 5.16 | 16
| 1 | 4 | 5.1741 | 5.1506 | 5.3146 | 4.7042 | 5.08 | 5.4 | 5.2 | 5.13 | 5.04 | 5.05 | 4.8876 | 5.04 | 5.026 | 5.25 | 4.89 | 3.81 | 5.07 | 5.02 | 5.0799
| 1 | 5 | 5.1123 | 51.1789 | 5.3585 | 4.7012 | 5.14 | 5.46 | 5.19 | 5.12 | 5.06 | 5.2 | 4.8881 | 5.19 | 5.0281 | 5.16 | 5.01 | 5.1 | 4.97 | 5.03 | 16
| 1 | 6 | 5.1159 | 5.1093 | 5.3363 | 4.8413 | 5.14 | 5.5 | 5.16 | 5.19 | 5.11 | 5 | 4.916 | 5.05 | 5.0311 | 5.12 | 5.11 | 5.05 | 4.95 | 5.11
| Mean |  | 5.1673 | 12.8134 | 5.32 | 4.7519 | 5.1033 | 5.4433 | 5.1717 | 5.105 | 5.1033 | 5.1433 | 4.8861 | 5.0517 | 5.0578 | 5.1667 | 5.0233 | 4.8583 | 4.9783 | 5.0617
| Median |  | 5.1666 | 5.1399 | 5.3255 | 4.7466 | 5.1 | 5.445 | 5.17 | 5.125 | 5.085 | 5.18 | 4.8879 | 5.045 | 5.0521 | 5.165 | 5.04 | 5.05 | 4.96 | 5.06
| Std Dev. |  | 0.048217 | 18.7952 | 0.030916 | 0.05181 | 0.034448 | 0.037238 | 0.022286 | 0.089163 | 0.077889 | 0.096471 | 0.031429 | 0.081343 | 0.033259 | 0.054283 | 0.093737 | 0.51425 | 0.050761 | 0.071949
| Rel.Std.Dev. |  | 0.9331% | 146.7% | 0.5811% | 1.09% | 0.675% | 0.6841% | 0.4309% | 1.747% | 1.526% | 1.876% | 0.6432% | 1.61% | 0.6576% | 1.051% | 1.866% | 10.58% | 1.02% | 1.421%
| PDM3 |  | 1.721% | 152.2% | 4.726% | -6.456% | 0.4612% | 7.154% | 1.806% | 0.494% | 0.4612% | 1.249% | -3.814% | -0.5559% | -0.4346% | 1.708% | -1.114% | -4.362% | -1.999% | -0.359%
| Z-Score (Absolute) |  | 0.65 | 78.56 | 2.21 | 3.58 | 0 | 3.46 | 0.7 | 0.02 | 0 | 0.41 | 2.21 | 0.53 | 0.46 | 0.65 | 0.82 | 2.5 | 1.27 | 0.42
| Analytical results for As in OREAS 254c (Certified Value 17.6 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 18 | 23.7365 | 152.8433 | 19 | 12.1 | 15.4 | 17.25 | 18.8 | 17 | 9.8686 | 16.9 | 16.95 | 17.5 | 18.6 | 19.8 | 18.8 | 15 |  | 1
| 1 | 2 | 18 | 23.9515 | 100.1715 | 19 | 13 | 15 | 15.69 | 18.8 | 17 | 9.7586 | 16.2 | 17.3084 | 17.3 | 18 | 19.2 | 18.3 | 17 |  | 21
| 1 | 3 | 17.8 | 24.1824 | 83.1902 | 19 | 10.6 | 15.5 | 16.4567 | 19.8 | 18 | 9.7841 | 17.2 | 17.601 | 17.5 | 18 | 19.5 | 17.2 | 16 |  | 16
| 1 | 4 | 17.9 | 23.2245 | 105.5761 | 19 | 11.6 | 15.2 | 15.9533 | 19 | 18 | 9.8241 | 16.3 | 15.1909 | 17.8 | 17.4 | 19.8 | 19.5 | 16 |  | 17.5969
| 1 | 5 | 17.6 | 23.6209 | 102.3739 | 19 | 9.9 | 15.6 | 17.4167 | 16.6 | 17 | 9.7805 | 17.6 | 17 | 17.7 | 17.7 | 19.3 | 18.8 | 17 |  | 17
| 1 | 6 | 17.8 | 24.2573 | 102.7187 | 19 | 11 | 15.6 | 16.5033 | 18.2 | 16 | 9.8312 | 17.6 | 15.0736 | 17.9 | 18.4 | 20.1 | 18.6 | 16
| Mean |  | 17.85 | 23.8289 | 107.8123 | 19 | 11.3667 | 15.3833 | 16.545 | 18.5333 | 17.1667 | 9.8079 | 16.9667 | 16.5207 | 17.6167 | 18.0167 | 19.6167 | 18.5333 | 16.1667
| Median |  | 17.85 | 23.844 | 102.5463 | 19 | 11.3 | 15.45 | 16.48 | 18.8 | 17 | 9.8041 | 17.05 | 16.975 | 17.6 | 18 | 19.65 | 18.7 | 16
| Std Dev. |  | 0.15166 | 0.38487 | 23.4646 | 0 | 1.1075 | 0.24014 | 0.6852 | 1.0783 | 0.75277 | 0.040589 | 0.61536 | 1.1013 | 0.22286 | 0.44008 | 0.34303 | 0.76333 | 0.75277
| Rel.Std.Dev. |  | 0.8496% | 1.615% | 21.76% | 0% | 9.744% | 1.561% | 4.141% | 5.818% | 4.385% | 0.4138% | 3.627% | 6.666% | 1.265% | 2.443% | 1.749% | 4.119% | 4.656%
| PDM3 |  | 1.438% | 35.42% | 512.7% | 7.974% | -35.41% | -12.58% | -5.978% | 5.322% | -2.445% | -44.26% | -3.581% | -6.116% | 0.1123% | 2.385% | 11.48% | 5.322% | -8.128%
| Z-Score (Absolute) |  | 0.07 | 3.42 | 50.6 | NA | 3.58 | 1.32 | 0.67 | 0.45 | NA | 4.45 | 0.43 | 0.68 | 0.07 | 0.16 | 1.06 | 0.45 | NA
Results from laboratories 1.07, 1.13 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ba in OREAS 254c (Certified Value 340 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 356.8 | 345.939 | 341.2283 | 319.9346 | 332 | 315 | 336 | 340 | 360 | 227 | 330.638 | 330 | 341.6713 | 335 | 350 | 360 | 330 | 362 | 1
| 1 | 2 | 351.1 | 350.548 | 336.7338 | 319.1698 | 336 | 327 | 344 | 341 | 350 | 223 | 331.6201 | 340 | 344.8614 | 328 | 350 | 360 | 330 | 382 | 22
| 1 | 3 | 347.2 | 346.082 | 342.3868 | 320.3153 | 336 | 321 | 347 | 345 | 340 | 226 | 331.0138 | 340 | 344.7667 | 339 | 340 | 360 | 320 | 376 | 16
| 1 | 4 | 351.2 | 344.3 | 336.344 | 316.3083 | 336 | 336 | 350 | 344 | 340 | 222 | 331.7453 | 340 | 343.552 | 340 | 340 | 290 | 340 | 369 | 339.6477
| 1 | 5 | 343.3 | 350.977 | 346.6974 | 318.0516 | 331 | 324 | 345 | 344 | 340 | 226 | 333.5708 | 350 | 335.9564 | 335 | 350 | 360 | 330 | 365 | 18
| 1 | 6 | 347.7 | 348.953 | 342.1451 | 327.6109 | 331 | 322 | 338 | 345 | 340 | 217 | 333.5434 | 340 | 335.2118 | 344 | 350 | 360 | 330 | 381
| Mean |  | 349.55 | 347.7998 | 340.9226 | 320.2317 | 333.6667 | 324.1667 | 343.3333 | 343.1667 | 345 | 223.5 | 332.0219 | 340 | 341.0033 | 336.8333 | 346.6667 | 348.3333 | 330 | 372.5
| Median |  | 349.4 | 347.5175 | 341.6867 | 319.5522 | 334 | 323 | 344.5 | 344 | 340 | 224.5 | 331.6827 | 340 | 342.6116 | 337 | 350 | 360 | 330 | 372.5
| Std Dev. |  | 4.5977 | 2.7437 | 3.8882 | 3.8941 | 2.582 | 7.0261 | 5.3541 | 2.137 | 8.3666 | 3.7283 | 1.2557 | 6.3246 | 4.359 | 5.4924 | 5.164 | 28.5774 | 6.3246 | 8.4083
| Rel.Std.Dev. |  | 1.315% | 0.7889% | 1.14% | 1.216% | 0.7738% | 2.167% | 1.559% | 0.6227% | 2.425% | 1.668% | 0.3782% | 1.86% | 1.278% | 1.631% | 1.49% | 8.204% | 1.917% | 2.257%
| PDM3 |  | 2.915% | 2.4% | 0.3754% | -5.716% | -1.761% | -4.558% | 1.085% | 1.036% | 1.576% | -34.2% | -2.245% | 0.1037% | 0.3991% | -0.8286% | 2.067% | 2.557% | -2.84% | 9.672%
| Z-Score (Absolute) |  | 0.82 | 0.65 | 0 | 1.98 | 0.7 | 1.6 | 0.23 | 0.21 | 0.39 | 11.21 | 0.85 | 0.09 | 0 | 0.39 | 0.54 | 0.7 | 1.05 | 3.01
| Analytical results for Be in OREAS 254c (Certified Value 1.93 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.84 | 2.024 | 1.4346 | 1.8 | 1.9 | 2 | < 5 | 1.98 | 2 | 1.6122 | 1.89 | 2.0462 | 1.9 | 1.94 | 1.94 | 1.81 | 2.15 |  | 1
| 1 | 2 | 1.83 | 1.9181 | 1.4357 | 1.7 | 2 | 2.02 | < 5 | 1.93 | 2 | 1.6185 | 1.93 | 2.0133 | 1.8 | 1.96 | 1.92 | 1.81 | 2.27 |  | 23
| 1 | 3 | 1.83 | 2.0141 | 1.426 | 1.8 | 2.1 | 1.99 | < 5 | 1.92 | 2 | 1.6002 | 1.91 | 2.0132 | 1.8 | 1.9 | 1.93 | 1.89 | 2.09 |  | 16
| 1 | 4 | 1.82 | 1.9949 | 1.4007 | 1.7 | 2 | 2.01 | < 5 | 1.9 | 2 | 1.6116 | 1.93 | 1.7498 | 1.9 | 1.86 | 1.48 | 1.84 | 2.23 |  | 1.9334
| 1 | 5 | 1.82 | 2.0274 | 1.457 | 1.8 | 2 | 2.02 | < 5 | 1.88 | 2 | 1.586 | 1.96 | 2.0092 | 2.1 | 1.93 | 1.94 | 1.82 | 2.21 |  | 19
| 1 | 6 | 1.85 | 2.0239 | 1.4446 | 1.8 | 1.9 | 1.98 | < 5 | 1.9 | 2 | 1.6142 | 1.93 | 1.7508 | 1.9 | 1.94 | 1.93 | 1.8 | 2.18
| Mean |  | 1.8317 | 2.0004 | 1.4331 | 1.7667 | 1.9833 | 2.0033 |  | 1.9183 | 2 | 1.6071 | 1.925 | 1.9304 | 1.9 | 1.9217 | 1.8567 | 1.8283 | 2.1883
| Median |  | 1.83 | 2.019 | 1.4352 | 1.8 | 2 | 2.005 |  | 1.91 | 2 | 1.6119 | 1.93 | 2.0112 | 1.9 | 1.935 | 1.93 | 1.815 | 2.195
| Std Dev. |  | 0.01169 | 0.04202 | 0.01904 | 0.05164 | 0.075277 | 0.01633 |  | 0.034881 | 0 | 0.011983 | 0.023452 | 0.14015 | 0.10954 | 0.036009 | 0.18468 | 0.033116 | 0.063377
| Rel.Std.Dev. |  | 0.6382% | 2.101% | 1.329% | 2.923% | 3.795% | 0.8151% |  | 1.818% | 0% | 0.7456% | 1.218% | 7.26% | 5.766% | 1.874% | 9.947% | 1.811% | 2.896%
| PDM3 |  | -5.26% | 3.468% | -25.87% | -8.622% | 2.585% | 3.62% |  | -0.777% | 3.447% | -16.87% | -0.4322% | -0.1519% | -1.725% | -0.6046% | -3.967% | -5.432% | 13.19%
| Z-Score (Absolute) |  | 0.73 | 0.66 | 4.01 | 1.26 | 0.52 | 0.69 | 4.78 | 0.01 | NA | 2.58 | 0.04 | 0.09 | 0.16 | 0.01 | 0.52 | 0.76 | 2.21
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 254c (Certified Value 0.35 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 0.36 |  | 0.34111 | < 0.1 | 0.35 | 0.38 | 0.24 | < 5 | 0.37 | 0.4 | 0.3605 | 0.32 | 0.32364 | 0.34 | 0.33 | 0.32 | 0.34 | 0.38 | 1
| 1 | 2 | 0.35 |  | 0.33003 | 1.7236 | 0.35 | 0.4 | 0.26 | < 5 | 0.37 | 0.4 | 0.3751 | 0.34 | 0.33892 | 0.34 | 0.31 | 0.31 | 0.35 | 0.39 | 24
| 1 | 3 | 0.34 |  | 0.34911 | 0.7049 | 0.35 | 0.39 | 0.26 | < 5 | 0.36 | 0.4 | 0.3533 | 0.35 | 0.33184 | 0.33 | 0.32 | 0.32 | 0.32 | 0.37 | 16
| 1 | 4 | 0.34 |  | 0.32201 | 0.674 | 0.35 | 0.34 | 0.25 | < 5 | 0.35 | 0.4 | 0.3532 | 0.32 | 0.31624 | 0.35 | 0.31 | 0.32 | 0.38 | 0.39 | 0.34686
| 1 | 5 | 0.34 |  | 0.34446 | 1.1348 | 0.34 | 0.37 | 0.27 | < 5 | 0.34 | 0.4 | 0.3569 | 0.34 | 0.31437 | 0.34 | 0.32 | 0.32 | 0.35 | 0.39 | 20
| 1 | 6 | 0.35 |  | 0.35049 | 0.546 | 0.36 | 0.39 | 0.25 | < 5 | 0.34 | 0.4 | 0.3692 | 0.34 | 0.34665 | 0.35 | 0.39 | 0.32 | 0.36 | 0.4
| Mean |  | 0.34667 |  | 0.33953 | 0.95666 | 0.35 | 0.37833 | 0.255 |  | 0.355 | 0.4 | 0.36137 | 0.335 | 0.32861 | 0.34167 | 0.33 | 0.31833 | 0.35 | 0.38667
| Median |  | 0.345 |  | 0.34278 | 0.7049 | 0.35 | 0.385 | 0.255 |  | 0.355 | 0.4 | 0.3587 | 0.34 | 0.32774 | 0.34 | 0.32 | 0.32 | 0.35 | 0.39
| Std Dev. |  | 0.008165 |  | 0.011278 | 0.48266 | 0.0063246 | 0.02137 | 0.010488 |  | 0.013784 | 6.0809e-17 | 0.0089714 | 0.012247 | 0.012828 | 0.0075277 | 0.030332 | 0.0040825 | 0.02 | 0.010328
| Rel.Std.Dev. |  | 2.355% |  | 3.322% | 50.45% | 1.807% | 5.648% | 4.113% |  | 3.883% | 1.52e-14% | 2.483% | 3.656% | 3.904% | 2.203% | 9.191% | 1.282% | 5.714% | 2.671%
| PDM3 |  | -0.0557% |  | -2.112% | 175.8% | 0.9053% | 9.074% | -26.48% |  | 2.347% | 15.32% | 4.182% | -3.419% | -5.261% | -1.497% | -4.861% | -8.224% | 0.9053% | 11.48%
| Z-Score (Absolute) |  | 0.07 | NA | 0.37 | 19.47 | 0.07 | 1.28 | 3.97 | 91.64 | 0.28 | NA | 0.56 | 0.57 | 0.84 | 0.28 | 0.78 | 1.28 | 0.07 | 1.63
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 254c (Certified Value 1.46 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.4817 | 1.4444 | 1.4704 | 2.1853 | 1.43 | 1.42 | 1.44 | 1.46 | 1.52 | 1.55 | 1.5157 | 1.49 | 1.4509 | 1.45 | 1.49 | 1.53 | 1.44 | 1.407 | 1
| 1 | 2 | 1.4865 | 1.4772 | 1.454 | 2.183 | 1.45 | 1.5 | 1.46 | 1.44 | 1.48 | 1.56 | 1.5151 | 1.53 | 1.4657 | 1.42 | 1.5 | 1.52 | 1.42 | 1.411 | #N/A
| 1 | 3 | 1.4661 | 1.4606 | 1.456 | 2.1513 | 1.44 | 1.49 | 1.46 | 1.47 | 1.48 | 1.55 | 1.5019 | 1.51 | 1.4494 | 1.44 | 1.44 | 1.51 | 1.43 | 1.415 | 16
| 1 | 4 | 1.4692 | 1.4498 | 1.4518 | 2.1361 | 1.43 | 1.38 | 1.45 | 1.46 | 1.46 | 1.55 | 1.5115 | 1.52 | 1.4598 | 1.45 | 1.42 | 1.08 | 1.46 | 1.381 | 1.4637
| 1 | 5 | 1.4559 | 1.4813 | 1.4724 | 2.1403 | 1.44 | 1.47 | 1.45 | 1.45 | 1.44 | 1.57 | 1.4958 | 1.55 | 1.4558 | 1.43 | 1.47 | 1.51 | 1.43 | 1.397 | 21
| 1 | 6 | 1.454 | 1.464 | 1.4742 | 2.1959 | 1.44 | 1.45 | 1.44 | 1.47 | 1.44 | 1.48 | 1.5148 | 1.52 | 1.4634 | 1.44 | 1.5 | 1.51 | 1.43 | 1.448
| Mean |  | 1.4689 | 1.4629 | 1.4631 | 2.1653 | 1.4383 | 1.4517 | 1.45 | 1.4583 | 1.47 | 1.5433 | 1.5091 | 1.52 | 1.4575 | 1.4383 | 1.47 | 1.4433 | 1.435 | 1.4098
| Median |  | 1.4676 | 1.4623 | 1.4632 | 2.1672 | 1.44 | 1.46 | 1.45 | 1.46 | 1.47 | 1.55 | 1.5132 | 1.52 | 1.4578 | 1.44 | 1.48 | 1.51 | 1.43 | 1.409
| Std Dev. |  | 0.01321 | 0.014569 | 0.010256 | 0.02577 | 0.0075277 | 0.04535 | 0.0089443 | 0.01169 | 0.030332 | 0.032042 | 0.0083216 | 0.02 | 0.0066337 | 0.01169 | 0.033466 | 0.17818 | 0.013784 | 0.022311
| Rel.Std.Dev. |  | 0.8993% | 0.996% | 0.701% | 1.19% | 0.5234% | 3.124% | 0.6168% | 0.8016% | 2.063% | 2.076% | 0.5514% | 1.316% | 0.4551% | 0.8128% | 2.277% | 12.34% | 0.9606% | 1.583%
| PDM3 |  | 0.356% | -0.05585% | -0.03865% | 47.94% | -1.732% | -0.8214% | -0.9353% | -0.3659% | 0.4312% | 5.441% | 3.106% | 3.847% | -0.4238% | -1.732% | 0.4312% | -1.391% | -1.96% | -3.679%
| Z-Score (Absolute) |  | 0.4 | 0.11 | 0.12 | 34.1 | 1.08 | 0.43 | 0.51 | 0.11 | 0.45 | 4 | 2.35 | 2.87 | 0.15 | 1.08 | 0.45 | 0.84 | 1.24 | 2.46
| Analytical results for Cd in OREAS 254c (Certified Value 0.071 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.07 | < 0.2 | 0.3472 | 0.08 | < 0.3 | 0.22 | < 0.5 | 0.05 | 0.2 | 0.0942 | 0.1 | 0.070336 | 0.31 | 0.06 | 0.07 | 0.05 |  |  | 1
| 1 | 2 | 0.08 | < 0.2 | 0.514 | 0.08 | < 0.3 | 0.23 | < 0.5 | 0.07 | 0.2 | 0.0899 | 0.07 | 0.068986 | 0.27 | 0.06 | 0.07 | 0.06 |  |  | 25
| 1 | 3 | 0.09 | < 0.2 | 0.3752 | 0.08 | < 0.3 | 0.23 | < 0.5 | 0.06 | 0.2 | 0.0913 | 0.05 | 0.086198 | 0.29 | 0.06 | 0.07 | 0.06 |  |  | 16
| 1 | 4 | 0.08 | < 0.2 | 0.5647 | 0.07 | < 0.3 | 0.24 | < 0.5 | 0.05 | 0.2 | 0.0995 | 0.06 | 0.085692 | 0.31 | 0.06 | 0.07 | 0.08 |  |  | 0.071267
| 1 | 5 | 0.06 | < 0.2 | 0.4291 | 0.09 | < 0.3 | 0.23 | < 0.5 | 0.06 | 0.2 | 0.0932 | 0.07 | 0.073922 | 0.3 | 0.06 | 0.07 | 0.04 |  |  | 22
| 1 | 6 | 0.09 | < 0.2 | 0.6139 | 0.06 | < 0.3 | 0.23 | < 0.5 | 0.06 | 0.2 | 0.1013 | 0.07 | 0.063871 | 0.3 | 0.08 | 0.06 | 0.05
| Mean |  | 0.078333 |  | 0.47402 | 0.076667 |  | 0.23 |  | 0.058333 | 0.2 | 0.0949 | 0.07 | 0.074834 | 0.29667 | 0.063333 | 0.068333 | 0.056667
| Median |  | 0.08 |  | 0.47155 | 0.08 |  | 0.23 |  | 0.06 | 0.2 | 0.0937 | 0.07 | 0.072129 | 0.3 | 0.06 | 0.07 | 0.055
| Std Dev. |  | 0.01169 |  | 0.10701 | 0.010328 |  | 0.0063246 |  | 0.0075277 | 3.0405e-17 | 0.0045488 | 0.016733 | 0.0091923 | 0.015055 | 0.008165 | 0.0040825 | 0.013663
| Rel.Std.Dev. |  | 14.92% |  | 22.57% | 13.47% |  | 2.75% |  | 12.9% | 1.52e-14% | 4.793% | 23.9% | 12.28% | 5.075% | 12.89% | 5.974% | 24.11%
| PDM3 |  | 9.916% |  | 565.1% | 7.577% |  | 222.7% |  | -18.15% | 180.6% | 33.16% | -1.777% | 5.006% | 316.3% | -11.13% | -4.116% | -20.49%
| Z-Score (Absolute) |  | 0.22 | 0.35 | 10.13 | 0.26 | 1.66 | 3.75 | 4.27 | 0.74 | 2.97 | 0.22 | 0.43 | 0.31 | 5.49 | 0.61 | 0.48 | 0.78
| Analytical results for Ce in OREAS 254c (Certified Value 70 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
1
| 1 | 1 | 69.84 | 70.903 | 66.4458 | 67.2 | 58 | 71.4 | 71 | 54.3302 | 69.5 | 71.3 | 69.4 | 70.9 | 67.1 | 69.81 |  |  |  |  | 1
| 1 | 2 | 68.82 | 71.6136 | 67.7694 | 69.3 | 57 | 71.2 | 67 | 54.7512 | 66 | 70.7 | 70.7 | 70.9 | 67.5 | 72.19 |  |  |  |  | 26
| 1 | 3 | 66.36 | 71.7741 | 74.686 | 67.4 | 59 | 69.8 | 71 | 54.0952 | 67.2 | 72.2 | 69.1 | 71.5 | 64.1 | 72.31 |  |  |  |  | 16
| 1 | 4 | 68.66 | 69.5748 | 66.0615 | 66.3 | 61 | 69.2 | 67 | 53.9292 | 64.5 | 71.7 | 68.6 | 70.7 | 72.6 | 70.45 |  |  |  |  | 69.5025
| 1 | 5 | 68.03 | 71.8381 | 68.9079 | 67.5 | 60 | 64.2 | 66 | 54.6017 | 73.3 | 71.1 | 69.1 | 70.3 | 70.6 | 72.34 |  |  |  |  | 23
| 1 | 6 | 66.46 | 71.3248 | 67.6495 | 67.5 | 59 | 69.3 | 69 | 54.1401 | 72.1 | 72 | 71.2 | 71.2 | 68.8 | 72.38
| Mean |  | 68.0283 | 71.1714 | 68.5867 | 67.5333 | 59 | 69.1833 | 68.5 | 54.3079 | 68.7667 | 71.5 | 69.6833 | 70.9167 | 68.45 | 71.58
| Median |  | 68.345 | 71.4692 | 67.7094 | 67.45 | 59 | 69.55 | 68 | 54.2351 | 68.35 | 71.5 | 69.25 | 70.9 | 68.15 | 72.25
| Std Dev. |  | 1.3821 | 0.85416 | 3.1563 | 0.97707 | 1.4142 | 2.6157 | 2.1679 | 0.31627 | 3.4789 | 0.56921 | 1.0265 | 0.41191 | 2.9535 | 1.143
| Rel.Std.Dev. |  | 2.032% | 1.2% | 4.602% | 1.447% | 2.397% | 3.781% | 3.165% | 0.5824% | 5.059% | 0.7961% | 1.473% | 0.5808% | 4.315% | 1.597%
| PDM3 |  | -2.121% | 2.401% | -1.318% | -2.833% | -15.11% | -0.4593% | -1.442% | -21.86% | -1.059% | 2.874% | 0.2601% | 2.035% | -1.514% | 2.989%
| Z-Score (Absolute) |  | 0.41 | 1.56 | 0.06 | 0.72 | 6.07 | 0.32 | 0.11 | 9.01 | 0.06 | 1.77 | 0.63 | 1.41 | 0.14 | 1.82
| Analytical results for Co in OREAS 254c (Certified Value 20.1 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 20.5 | 21.348 | 20.4751 | 21.1526 | 19 | 20 | 20.6 | 22.4647 | 20 | 19.4 | 17.8927 | 18.3 | 19.3756 | 20.32 | 20.8 | 21.7 | 18.8 | 20.5 | 1
| 1 | 2 | 20.2 | 21.132 | 20.2905 | 20.4203 | 19 | 20 | 20.4 | 22.4311 | 19.5 | 19.5 | 18.0637 | 17.7 | 19.7851 | 19.81 | 21.1 | 20.7 | 19 | 21.2 | 27
| 1 | 3 | 19.9 | 20.25 | 20.3921 | 21.3262 | 19 | 20 | 20.6 | 22.6517 | 19.6 | 19.8 | 17.7406 | 17.7 | 19.6071 | 20.31 | 20.3 | 21.7 | 20.1 | 21.6 | 16
| 1 | 4 | 19.8 | 20.376 | 20.0659 | 20.8576 | 19 | 20 | 20.7 | 22.651 | 19.4 | 19.2 | 17.7804 | 17.5 | 19.0674 | 20.66 | 20 | 21.3 | 19.8 | 21 | 20.097
| 1 | 5 | 19.7 | 20.718 | 20.5076 | 21.0945 | 19 | 20 | 20.5 | 22.9457 | 19.8 | 19.9 | 17.6431 | 20.3 | 19.4018 | 20.13 | 20.2 | 20.9 | 19.5 | 20.7 | 24
| 1 | 6 | 20.1 | 20.934 | 20.5496 | 21.3079 | 19 | 20 | 20.5 | 23.1165 | 20.3 | 19 | 17.902 | 18.8 | 19.8056 | 20.72 | 21.3 | 21.2 | 18.9 | 21.8
| Mean |  | 20.0333 | 20.793 | 20.3801 | 21.0265 | 19 | 20 | 20.55 | 22.7101 | 19.7667 | 19.4667 | 17.8371 | 18.3833 | 19.5071 | 20.325 | 20.6167 | 21.25 | 19.35 | 21.1333
| Median |  | 20 | 20.826 | 20.4336 | 21.1236 | 19 | 20 | 20.55 | 22.6513 | 19.7 | 19.45 | 17.8366 | 18 | 19.5044 | 20.315 | 20.55 | 21.25 | 19.25 | 21.1
| Std Dev. |  | 0.29439 | 0.4283 | 0.17939 | 0.34221 | 0 | 0 | 0.10488 | 0.27041 | 0.33862 | 0.34448 | 0.14753 | 1.0553 | 0.28217 | 0.33816 | 0.52694 | 0.40866 | 0.53198 | 0.50465
| Rel.Std.Dev. |  | 1.47% | 2.06% | 0.8802% | 1.628% | 0% | 0% | 0.5104% | 1.191% | 1.713% | 1.77% | 0.8271% | 5.741% | 1.446% | 1.664% | 2.556% | 1.923% | 2.749% | 2.388%
| PDM3 |  | -0.3167% | 3.463% | 1.409% | 4.625% | -5.458% | -0.4825% | 2.254% | 13% | -1.644% | -3.136% | -11.24% | -8.527% | -2.935% | 1.135% | 2.586% | 5.737% | -3.717% | 5.157%
| Z-Score (Absolute) |  | 0.14 | 0.6 | 0.2 | 0.83 | 1.15 | 0.17 | 0.36 | 2.47 | 0.4 | 0.69 | 2.28 | 1.75 | 0.65 | 0.14 | 0.43 | 1.04 | 0.81 | 0.93
| Analytical results for Cr in OREAS 254c (Certified Value 111 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 126 | 112.0492 | 137.6008 | 104 | 124 | 119 | 100.2986 | 117 | 98 | 91.4376 | 113 | 111.6485 | 112 | 115 | 112 | 106 |  |  | 1
| 1 | 2 | 123 | 107.8205 | 130.9978 | 108 | 99 | 114 | 100.0498 | 112 | 105 | 95.2342 | 119 | 111.0471 | 109 | 112 | 115 | 106 |  |  | 28
| 1 | 3 | 129 | 112.6174 | 140.4524 | 103 | 94 | 115 | 103.818 | 110 | 101 | 91.2389 | 116 | 109.0345 | 114 | 113 | 113 | 111 |  |  | 16
| 1 | 4 | 116 | 111.4441 | 138.2764 | 111 | 98 | 113 | 104.1288 | 111 | 106 | 92.8357 | 117 | 109.0149 | 117 | 114 | 99 | 104 |  |  | 110.56
| 1 | 5 | 126 | 111.466 | 136.2573 | 117 | 92 | 119 | 102.1484 | 112 | 106 | 91.04 | 118 | 111.6523 | 112 | 127 | 114 | 104 |  |  | 25
| 1 | 6 | 129 | 113.2935 | 130.4428 | 111 | 106 | 115 | 104.2176 | 113 | 94 | 94.7325 | 117 | 109.2893 | 115 | 113 | 116 | 104
| Mean |  | 124.8333 | 111.4484 | 135.6713 | 109 | 102.1667 | 115.8333 | 102.4435 | 112.5 | 101.6667 | 92.7532 | 116.6667 | 110.2811 | 113.1667 | 115.6667 | 111.5 | 105.8333
| Median |  | 126 | 111.7576 | 136.929 | 109.5 | 98.5 | 115 | 102.9832 | 112 | 103 | 92.1367 | 117 | 110.1682 | 113 | 113.5 | 113.5 | 105
| Std Dev. |  | 4.8751 | 1.9128 | 4.0716 | 5.1769 | 11.7374 | 2.5626 | 1.9128 | 2.429 | 4.9261 | 1.846 | 2.0656 | 1.3021 | 2.7869 | 5.6451 | 6.2849 | 2.7142
| Rel.Std.Dev. |  | 3.905% | 1.716% | 3.001% | 4.749% | 11.49% | 2.212% | 1.867% | 2.159% | 4.845% | 1.99% | 1.771% | 1.181% | 2.463% | 4.88% | 5.637% | 2.565%
| PDM3 |  | 12.91% | 0.8036% | 22.71% | -1.411% | -7.592% | 4.77% | -7.341% | 1.755% | -8.044% | -16.11% | 5.523% | -0.2522% | 2.358% | 4.619% | 0.8502% | -4.275%
| Z-Score (Absolute) |  | 1.89 | 0 | 3.42 | 0.35 | 1.31 | 0.62 | 1.28 | 0.14 | 1.38 | 2.64 | 0.73 | 0.17 | 0.24 | 0.59 | 0 | 0.8
| Analytical results for Cs in OREAS 254c (Certified Value 2.47 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.55 | 2.7309 | 2.3 | 2.4 | 2.51 | 2.61 | 2.3919 | 2.42 | 2.4484 | 2.38 | 2.39 | 2.57 | 2.41 | 2.58 |  |  |  |  | 1
| 1 | 2 | 2.52 | 2.64 | 2.4 | 2.46 | 2.61 | 2.52 | 2.3577 | 2.4 | 2.5029 | 2.33 | 2.37 | 2.54 | 2.44 | 2.72 |  |  |  |  | 29
| 1 | 3 | 2.47 | 2.5887 | 2.3 | 2.38 | 2.59 | 2.53 | 2.3968 | 2.36 | 2.4534 | 2.43 | 2.4 | 2.62 | 2.23 | 2.67 |  |  |  |  | 16
| 1 | 4 | 2.53 | 2.5571 | 2.3 | 2.43 | 2.64 | 2.54 | 2.2868 | 2.36 | 2.1974 | 2.43 | 2.34 | 2.61 | 2.55 | 2.67 |  |  |  |  | 2.4745
| 1 | 5 | 2.47 | 2.5939 | 2.3 | 2.39 | 2.64 | 2.29 | 2.3645 | 2.6 | 2.5064 | 2.4 | 2.39 | 2.51 | 2.51 | 2.72 |  |  |  |  | 26
| 1 | 6 | 2.46 | 2.5739 | 2.4 | 2.45 | 2.54 | 2.48 | 2.4095 | 2.59 | 2.2036 | 2.44 | 2.44 | 2.58 | 2.51 | 2.63
| Mean |  | 2.5 | 2.6141 | 2.3333 | 2.4183 | 2.5883 | 2.495 | 2.3679 | 2.455 | 2.3854 | 2.4017 | 2.3883 | 2.5717 | 2.4417 | 2.665
| Median |  | 2.495 | 2.5913 | 2.3 | 2.415 | 2.6 | 2.525 | 2.3782 | 2.41 | 2.4509 | 2.415 | 2.39 | 2.575 | 2.475 | 2.67
| Std Dev. |  | 0.037947 | 0.06363 | 0.05164 | 0.033116 | 0.053448 | 0.10895 | 0.04435 | 0.11095 | 0.14521 | 0.041673 | 0.033116 | 0.041673 | 0.11566 | 0.053944
| Rel.Std.Dev. |  | 1.518% | 2.434% | 2.213% | 1.369% | 2.065% | 4.367% | 1.873% | 4.519% | 6.088% | 1.735% | 1.387% | 1.62% | 4.737% | 2.024%
| PDM3 |  | 1.03% | 5.64% | -5.706% | -2.271% | 4.599% | 0.8277% | -4.31% | -0.7888% | -3.603% | -2.944% | -3.483% | 3.926% | -1.328% | 7.698%
| Z-Score (Absolute) |  | 0.57 | 1.82 | 1.26 | 0.33 | 1.54 | 0.51 | 0.88 | 0.07 | 0.69 | 0.51 | 0.66 | 1.35 | 0.07 | 2.38
| Analytical results for Cu in OREAS 254c (Certified Value 31.4 ppm) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 31.7 | 31.5314 | 29.9936 | 28 | 35 | 36 | 29.9 | 31.5 | 35.1 | 19.0441 | 29 | 25.4479 | 32 | 31 | 34.1 | 31.2 | 32.7 |  | 1
| 1 | 2 | 32 | 31.0603 | 28.8921 | 28 | 32 | 36 | 28.9275 | 31.2 | 35.2 | 18.7771 | 28.2 | 24.9846 | 32.3 | 31.6 | 33.5 | 31.8 | 33.6 |  | 30
| 1 | 3 | 32.2 | 30.886 | 29.726 | 28 | 32 | 36 | 30.0248 | 31.5 | 35.3 | 19.5529 | 30.2 | 26.1888 | 33.3 | 30.9 | 35 | 31.2 | 33.6 |  | 16
| 1 | 4 | 32.1 | 30.9728 | 28.7408 | 28 | 32 | 36 | 30.4833 | 30.7 | 34.7 | 18.7759 | 27.6 | 24.8216 | 33 | 30.2 | 33 | 33.4 | 32 |  | 31.3785
| 1 | 5 | 31.9 | 31.4089 | 29.17 | 28 | 33 | 36 | 31.0958 | 30 | 35.7 | 18.9305 | 32.1 | 26.0724 | 33.3 | 30.6 | 33.2 | 32.5 | 32.1 |  | 27
| 1 | 6 | 32.2 | 31.6393 | 29.5712 | 28 | 32 | 36 | 28.9833 | 30.8 | 34.3 | 19.6757 | 30.1 | 24.5097 | 32.8 | 31.8 | 33.6 | 31.8 | 33.7
| Mean |  | 32.0167 | 31.2498 | 29.3489 | 28 | 32.6667 | 36 | 29.9025 | 30.95 | 35.05 | 19.126 | 29.5333 | 25.3375 | 32.7833 | 31.0167 | 33.7333 | 31.9833 | 32.95
| Median |  | 32.05 | 31.2346 | 29.3706 | 28 | 32 | 36 | 29.9624 | 31 | 35.15 | 18.9873 | 29.55 | 25.2163 | 32.9 | 30.95 | 33.55 | 31.8 | 33.15
| Std Dev. |  | 0.19408 | 0.31665 | 0.4935 | 0 | 1.2111 | 0 | 0.84518 | 0.57533 | 0.48888 | 0.39336 | 1.6219 | 0.686 | 0.53448 | 0.60139 | 0.72572 | 0.84479 | 0.78677
| Rel.Std.Dev. |  | 0.6062% | 1.013% | 1.681% | 0% | 3.707% | 0% | 2.826% | 1.859% | 1.395% | 2.057% | 5.492% | 2.707% | 1.63% | 1.939% | 2.151% | 2.641% | 2.388%
| PDM3 |  | 2.034% | -0.4101% | -6.468% | -10.77% | 4.105% | 14.73% | -4.704% | -1.365% | 11.7% | -39.05% | -5.88% | -19.25% | 4.477% | -1.153% | 7.505% | 1.928% | 5.008%
| Z-Score (Absolute) |  | 0.3 | 0 | 0.75 | 1.29 | 0.56 | 1.88 | 0.53 | 0.12 | 1.51 | 4.81 | 0.68 | 2.34 | 0.61 | 0.09 | 0.98 | 0.29 | 0.67
| Analytical results for Dy in OREAS 254c (Certified Value 3.44 ppm) | [19 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 3.39 | 3.5255 | 4.2354 | 3.5 | 3.43 | 2.6293 | 3.18 | [11 empty] | 1
| 1 | 2 | 3.4 | 3.5548 | 4.0328 | 3.6 | 3.42 | 2.6634 | 3.37 | [11 empty] | 4
| 1 | 3 | 3.38 | 3.5812 | 3.8208 | 3.3 | 3.49 | 2.5917 | 3.4 | [11 empty] | 16
| 1 | 4 | 3.37 | 3.5489 | 3.8607 | 3.6 | 3.52 | 2.6248 | 3.32 | [11 empty] | 3.444
| 1 | 5 | 3.36 | 3.5334 | 3.9222 | 3.5 | 3.5 | 2.6303 | 3.34 | [11 empty] | 28
| 1 | 6 | 3.34 | 3.6156 | 3.9139 | 3.5 | 3.45 | 2.6531 | 3.3
| Mean |  | 3.3733 | 3.5599 | 3.9643 | 3.5 | 3.4683 | 2.6321 | 3.3183
| Median |  | 3.375 | 3.5518 | 3.918 | 3.5 | 3.47 | 2.6298 | 3.33
| Std Dev. |  | 0.021602 | 0.033426 | 0.15089 | 0.10954 | 0.040702 | 0.024991 | 0.076529
| Rel.Std.Dev. |  | 0.6404% | 0.939% | 3.806% | 3.13% | 1.174% | 0.9494% | 2.306%
| PDM3 |  | -2.051% | 3.366% | 15.11% | 1.627% | 0.7071% | -23.57% | -3.648%
| Z-Score (Absolute) |  | 0.67 | 0.65 | 3.52 | 0.22 | 0 | 5.94 | 1.06
| Analytical results for Er in OREAS 254c (Certified Value 1.59 ppm) | [19 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.59 | 1.5873 | 1.6274 | 1.6 | 1.55 | 3.0332 | 1.59 | [11 empty] | 1
| 1 | 2 | 1.64 | 1.6522 | 1.5922 | 1.7 | 1.51 | 3.0628 | 1.68 | [11 empty] | 5
| 1 | 3 | 1.51 | 1.6597 | 1.5355 | 1.6 | 1.55 | 2.9952 | 1.65 | [11 empty] | 16
| 1 | 4 | 1.56 | 1.6431 | 1.517 | 1.6 | 1.6 | 3.0099 | 1.62 | [11 empty] | 1.5928
| 1 | 5 | 1.54 | 1.69 | 1.524 | 1.6 | 1.57 | 2.9769 | 1.65 | [11 empty] | 29
| 1 | 6 | 1.47 | 1.6468 | 1.545 | 1.6 | 1.54 | 3.046 | 1.6
| Mean |  | 1.5517 | 1.6465 | 1.5569 | 1.6167 | 1.5533 | 3.0206 | 1.6317
| Median |  | 1.55 | 1.6495 | 1.5402 | 1.6 | 1.55 | 3.0215 | 1.635
| Std Dev. |  | 0.059805 | 0.033485 | 0.043575 | 0.040825 | 0.030111 | 0.032424 | 0.034303
| Rel.Std.Dev. |  | 3.854% | 2.034% | 2.799% | 2.525% | 1.938% | 1.073% | 2.102%
| PDM3 |  | -2.581% | 3.373% | -2.256% | 1.499% | -2.477% | 89.65% | 2.441%
| Z-Score (Absolute) |  | 0.73 | 0.34 | 0.67 | 0 | 0.71 | 15.83 | 0.17
| Analytical results for Eu in OREAS 254c (Certified Value 1.5 ppm) | [19 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.58 | 0.5028 | 1.6474 | 1.33 | 1.57 | 1.3244 | 1.51 | [11 empty] | 1
| 1 | 2 | 1.59 | 0.46186 | 1.6753 | 1.4 | 1.56 | 1.3351 | 1.6 | [11 empty] | 6
| 1 | 3 | 1.58 | 0.50909 | 1.5899 | 1.34 | 1.59 | 1.3162 | 1.6 | [11 empty] | 16
| 1 | 4 | 1.58 | 0.47035 | 1.525 | 1.41 | 1.61 | 1.3359 | 1.53 | [11 empty] | 1.5004
| 1 | 5 | 1.54 | 0.49987 | 1.5763 | 1.34 | 1.59 | 1.3179 | 1.53 | [11 empty] | 30
| 1 | 6 | 1.57 | 0.49728 | 1.5668 | 1.35 | 1.56 | 1.3453 | 1.6
| Mean |  | 1.5733 | 0.49021 | 1.5968 | 1.3617 | 1.58 | 1.3291 | 1.5617
| Median |  | 1.58 | 0.49857 | 1.5831 | 1.345 | 1.58 | 1.3298 | 1.565
| Std Dev. |  | 0.017512 | 0.01927 | 0.055235 | 0.034303 | 0.02 | 0.011467 | 0.042622
| Rel.Std.Dev. |  | 1.113% | 3.931% | 3.459% | 2.519% | 1.266% | 0.8628% | 2.729%
| PDM3 |  | 4.859% | -67.33% | 6.422% | -9.248% | 5.303% | -11.42% | 4.081%
| Z-Score (Absolute) |  | 0.22 | 20.57 | 0.67 | 3.84 | 0.35 | 4.46 | 0
| Analytical results for Fe in OREAS 254c (Certified Value 3.93 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.05 | 4.1769 | 4.029 | 3.1699 | 4.04 | 3.66 | 3.9 | 3.93 | 4.13 | 3.82 | 3.9781 | 3.89 | 3.9527 | 3.92 | 3.94 | 4.04 | 3.8 | 3.89 | 1
| 1 | 2 | 4.04 | 4.1935 | 3.9787 | 3.1487 | 4.03 | 3.79 | 3.87 | 3.87 | 4.03 | 3.83 | 3.9279 | 4 | 3.9946 | 3.85 | 3.98 | 3.99 | 3.83 | 3.81 | #N/A
| 1 | 3 | 3.99 | 4.1481 | 3.9974 | 3.2101 | 3.99 | 3.8 | 3.92 | 3.95 | 3.98 | 3.86 | 3.9309 | 3.94 | 3.9924 | 3.87 | 3.86 | 4.02 | 3.85 | 3.99 | 16
| 1 | 4 | 4 | 4.1283 | 3.9775 | 3.1632 | 4 | 4.03 | 4.03 | 3.95 | 3.94 | 3.81 | 3.9363 | 3.95 | 3.9512 | 3.95 | 3.82 | 3.05 | 3.9 | 3.92 | 3.9332
| 1 | 5 | 3.94 | 4.1485 | 4.0102 | 3.1539 | 4.07 | 3.71 | 4.01 | 3.89 | 3.9 | 3.91 | 3.9135 | 4.04 | 3.9645 | 3.88 | 3.9 | 4.05 | 3.82 | 3.78 | 31
| 1 | 6 | 3.96 | 4.1875 | 4.0271 | 3.2264 | 4.04 | 3.68 | 3.9 | 3.97 | 3.95 | 3.72 | 3.922 | 3.99 | 3.9723 | 3.89 | 3.96 | 4 | 3.82 | 3.94
| Mean |  | 3.9967 | 4.1638 | 4.0033 | 3.1787 | 4.0283 | 3.7783 | 3.9383 | 3.9267 | 3.9883 | 3.825 | 3.9348 | 3.9683 | 3.9713 | 3.8933 | 3.91 | 3.8583 | 3.8367 | 3.8883
| Median |  | 3.995 | 4.1627 | 4.0038 | 3.1665 | 4.035 | 3.75 | 3.91 | 3.94 | 3.965 | 3.825 | 3.9294 | 3.97 | 3.9684 | 3.885 | 3.92 | 4.01 | 3.825 | 3.905
| Std Dev. |  | 0.043205 | 0.025918 | 0.022712 | 0.031904 | 0.029269 | 0.13586 | 0.065549 | 0.038816 | 0.081833 | 0.062849 | 0.022605 | 0.052694 | 0.018891 | 0.036148 | 0.061644 | 0.39666 | 0.035024 | 0.079854
| Rel.Std.Dev. |  | 1.081% | 0.6225% | 0.5673% | 1.004% | 0.7266% | 3.596% | 1.664% | 0.9885% | 2.052% | 1.643% | 0.5745% | 1.328% | 0.4757% | 0.9285% | 1.577% | 10.28% | 0.9129% | 2.054%
| PDM3 |  | 1.613% | 5.863% | 1.783% | -19.18% | 2.418% | -3.938% | 0.1302% | -0.1664% | 1.401% | -2.751% | 0.03925% | 0.893% | 0.9681% | -1.014% | -0.5901% | -1.904% | -2.455% | -1.141%
| Z-Score (Absolute) |  | 0.72 | 2.54 | 0.79 | 8.21 | 1.07 | 1.66 | 0.08 | 0.04 | 0.63 | 1.15 | 0.04 | 0.41 | 0.44 | 0.41 | 0.23 | 0.79 | 1.03 | 0.46
| Analytical results for Ga in OREAS 254c (Certified Value 15 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 15.42 | 14.44 | 14.8183 | 16.3 | 9.9 | 16.64 | 16.4552 | 15.3 | 13.0045 | 13.65 | 15.3543 | 15.69 | 13.95 | 15.15 | 14.35 | 15.63 |  |  | 1
| 1 | 2 | 15.29 | 13.18 | 14.4127 | 16.2 | 10.3 | 16.85 | 16.1795 | 14.75 | 12.6377 | 13.25 | 15.3693 | 15.47 | 13.85 | 14.7 | 14.5 | 16.21 |  |  | 34
| 1 | 3 | 15.27 | 13.36 | 14.9623 | 16.2 | 10.7 | 16.92 | 16.6757 | 15.05 | 12.4569 | 13.5 | 14.837 | 15.89 | 14 | 15.4 | 14.2 | 16.32 |  |  | 16
| 1 | 4 | 15.52 | 13.5 | 15.06 | 16.5 | 14.8 | 16.92 | 16.6582 | 15.15 | 12.5446 | 13.05 | 14.2408 | 16.3 | 13.5 | 15.1 | 15.55 | 15.78 |  |  | 15.037
| 1 | 5 | 15.23 | 13.48 | 15.2054 | 16.5 | 9.6 | 17.28 | 16.7289 | 14.75 | 12.4928 | 15.25 | 14.4372 | 15.73 | 13.8 | 14.75 | 14.8 | 16.54 |  |  | 32
| 1 | 6 | 15.47 | 14.46 | 14.6631 | 16.5 | 9.3 | 16.55 | 16.9985 | 15.55 | 12.5413 | 14.3 | 14.6676 | 16.02 | 14.2 | 15.2 | 14.45 | 16.93
| Mean |  | 15.3667 | 13.7367 | 14.8536 | 16.3667 | 10.7667 | 16.86 | 16.616 | 15.0917 | 12.613 | 13.8333 | 14.8177 | 15.85 | 13.8833 | 15.05 | 14.6417 | 16.235
| Median |  | 15.355 | 13.49 | 14.8903 | 16.4 | 10.1 | 16.885 | 16.6669 | 15.1 | 12.543 | 13.575 | 14.7523 | 15.81 | 13.9 | 15.125 | 14.475 | 16.265
| Std Dev. |  | 0.11911 | 0.56419 | 0.28641 | 0.15055 | 2.0373 | 0.2562 | 0.27596 | 0.31371 | 0.20124 | 0.81527 | 0.46747 | 0.28893 | 0.23381 | 0.27203 | 0.48725 | 0.48103
| Rel.Std.Dev. |  | 0.7751% | 4.107% | 1.928% | 0.9199% | 18.92% | 1.52% | 1.661% | 2.079% | 1.596% | 5.894% | 3.155% | 1.823% | 1.684% | 1.808% | 3.328% | 2.963%
| PDM3 |  | 2.192% | -8.648% | -1.22% | 8.842% | -28.4% | 12.12% | 10.5% | 0.3634% | -16.12% | -8.005% | -1.459% | 5.407% | -7.672% | 0.08632% | -2.629% | 7.967%
| Z-Score (Absolute) |  | 0.26 | 0.75 | 0.06 | 0.87 | 2.58 | 1.18 | 1.03 | 0.09 | 1.44 | 0.69 | 0.08 | 0.55 | 0.66 | 0.06 | 0.19 | 0.79
| Analytical results for Gd in OREAS 254c (Certified Value 5.31 ppm) | [19 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 4.94 | 5.9702 | 5.6899 | 4.9 | 4.95 | 2.7168 | 5.78 | [11 empty] | 1
| 1 | 2 | 4.87 | 5.7843 | 5.3747 | 5.3 | 4.99 | 2.8145 | 5.9 | [11 empty] | 7
| 1 | 3 | 4.81 | 5.9373 | 5.2165 | 4.9 | 5 | 2.7082 | 5.92 | [11 empty] | 16
| 1 | 4 | 4.92 | 6.0457 | 5.4131 | 5 | 4.9 | 2.8009 | 5.87 | [11 empty] | 5.3145
| 1 | 5 | 4.88 | 5.8073 | 5.2092 | 4.8 | 4.97 | 2.7594 | 5.71 | [11 empty] | 33
| 1 | 6 | 4.72 | 5.923 | 5.2394 | 4.7 | 4.88 | 2.8183 | 6.1
| Mean |  | 4.8567 | 5.9113 | 5.3571 | 4.9333 | 4.9483 | 2.7697 | 5.88
| Median |  | 4.875 | 5.9302 | 5.3071 | 4.9 | 4.96 | 2.7801 | 5.885
| Std Dev. |  | 0.080664 | 0.099284 | 0.18421 | 0.20656 | 0.048751 | 0.049051 | 0.13372
| Rel.Std.Dev. |  | 1.661% | 1.68% | 3.439% | 4.187% | 0.9852% | 1.771% | 2.274%
| PDM3 |  | -8.614% | 11.23% | 0.8028% | -7.171% | -6.889% | -47.88% | 10.64%
| Z-Score (Absolute) |  | 0.15 | 1.59 | 0.67 | 0.02 | 0 | 3.59 | 1.54
| Analytical results for Ge in OREAS 254c (Indicative Value 0.13 ppm) | [19 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.2 | < 0.2 | 1.6 | < 0.1 | 0.09 | 0.2 | 0.11 | 0.0754 | 0.09 | 0.15421 | 0.15 | 0.23 | 0.1 | [5 empty] | 1
| 1 | 2 | 1.3 | < 0.2 | 1.5 | < 0.1 | 0.09 | 0.22 | 0.06 | 0.0482 | 0.1 | 0.19087 | 0.14 | 0.23 | 0.11 | [5 empty] | 1
| 1 | 3 | 1.2 | < 0.2 | 1.4 | < 0.1 | 0.08 | 0.22 | 0.17 | 0.1725 | 0.08 | 0.19092 | 0.12 | 0.23 | 0.16 | [5 empty] | 16
| 1 | 4 | 1.2 | < 0.2 | 1.6 | < 0.1 | 0.09 | 0.24 | < 0.05 | 0.0595 | 0.08 | 0.12908 | 0.12 | 0.23 | 0.09 | [5 empty] | 0.13107
| 1 | 5 | 1.1 | < 0.2 | 1.5 | < 0.1 | 0.09 | 0.12 | 0.08 | 0.2727 | 0.1 | 0.17667 | 0.13 | 0.21 | 0.11 | [5 empty] | 7
| 1 | 6 | 1.3 | < 0.2 | 1.5 | < 0.1 | 0.1 | 0.15 | < 0.05 | 0.0469 | 0.08 | 0.13316 | 0.15 | 0.23 | 0.09
| Mean |  | 1.2167 |  | 1.5167 |  | 0.09 | 0.19167 | 0.105 | 0.11253 | 0.088333 | 0.16249 | 0.135 | 0.22667 | 0.11
| Median |  | 1.2 |  | 1.5 |  | 0.09 | 0.21 | 0.095 | 0.06745 | 0.085 | 0.16544 | 0.135 | 0.23 | 0.105
| Std Dev. |  | 0.075277 |  | 0.075277 |  | 0.0063246 | 0.046655 | 0.047958 | 0.091528 | 0.0098319 | 0.027784 | 0.013784 | 0.008165 | 0.026077
| Rel.Std.Dev. |  | 6.187% |  | 4.963% |  | 7.027% | 24.34% | 45.67% | 81.33% | 11.13% | 17.1% | 10.21% | 3.602% | 23.71%
| PDM3 |  | 828.2% |  | 1057% |  | -31.34% | 46.23% | -19.89% | -14.14% | -32.61% | 23.97% | 2.996% | 72.93% | -16.08%
| Z-Score (Absolute) |  | 21.77 | 0.25 | 27.68 | 1.23 | 0.44 | 1.56 | 0.67 | 0 | 0.48 | 0.98 | 0.44 | 2.25 | 0.05
| Analytical results for Hf in OREAS 254c (Certified Value 4.41 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 4.74 | 5.0525 | 4.6946 | 4.3 | < 0.1 | 3.8 | 4.7 | 4.3 | 4.9006 | 4.1 | 4.1993 | 4.11 | 4.3 | 4.7 | 4.3 | 4.27 |  |  | 1
| 1 | 2 | 4.56 | 4.9803 | 4.5901 | 4.2 | < 0.1 | 3.8 | 4.5 | 4.2 | 4.9182 | 4 | 4.4348 | 4.42 | 4.4 | 4.6 | 4.2 | 4.28 |  |  | 8
| 1 | 3 | 4.55 | 5.1591 | 4.4667 | 4.3 | 0.3 | 3.9 | 4.5 | 4.4 | 4.7977 | 4 | 4.2466 | 4.14 | 4.2 | 5 | 4.4 | 4.35 |  |  | 16
| 1 | 4 | 4.53 | 4.987 | 4.4739 | 4.3 | 0.1 | 4.1 | 4.4 | 4.2 | 4.7329 | 3.9 | 3.7131 | 4.19 | 4.1 | 4.8 | 4.5 | 4.19 |  |  | 4.4058
| 1 | 5 | 4.62 | 5.0089 | 4.5218 | 4.3 | 0.3 | 3.9 | 4.5 | 4.5 | 4.8859 | 4.6 | 4.2405 | 4.2 | 4.2 | 4.7 | 4.5 | 4.2 |  |  | 34
| 1 | 6 | 4.47 | 4.982 | 4.4137 | 4.3 | 0.2 | 4 | 4.6 | 4.5 | 5.0871 | 4.2 | 3.6878 | 4.16 | 4.5 | 4.9 | 4.4 | 4.34
| Mean |  | 4.5783 | 5.0283 | 4.5268 | 4.2833 | 0.225 | 3.9167 | 4.5333 | 4.35 | 4.8871 | 4.1333 | 4.087 | 4.2033 | 4.2833 | 4.7833 | 4.3833 | 4.2717
| Median |  | 4.555 | 4.9979 | 4.4979 | 4.3 | 0.25 | 3.9 | 4.5 | 4.35 | 4.8933 | 4.05 | 4.2199 | 4.175 | 4.25 | 4.75 | 4.4 | 4.275
| Std Dev. |  | 0.092826 | 0.069586 | 0.10137 | 0.040825 | 0.095743 | 0.1169 | 0.10328 | 0.13784 | 0.12076 | 0.25033 | 0.31041 | 0.11112 | 0.1472 | 0.1472 | 0.1169 | 0.067355
| Rel.Std.Dev. |  | 2.028% | 1.384% | 2.239% | 0.9531% | 42.55% | 2.985% | 2.278% | 3.169% | 2.471% | 6.056% | 7.595% | 2.644% | 3.436% | 3.077% | 2.667% | 1.577%
| PDM3 |  | 3.917% | 14.13% | 2.747% | -2.779% | -94.89% | -11.1% | 2.896% | -1.266% | 10.92% | -6.183% | -7.234% | -4.595% | -2.779% | 8.57% | -0.509% | -3.044%
| Z-Score (Absolute) |  | 0.83 | 2.25 | 0.66 | 0.11 | 13.11 | 1.26 | 0.68 | 0.11 | 1.8 | 0.58 | 0.73 | 0.36 | 0.11 | 1.47 | 0.21 | 0.14
| Analytical results for Hg in OREAS 254c (Indicative Value 0.058 ppm) | [19 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.10 | 1.14 | [15 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [15 empty] | 3
3
| 1 | 1 | < 1 | < 2 | 0.058 | [15 empty] | 1
| 1 | 2 | < 1 | < 2 | 0.0576 | [15 empty] | 2
| 1 | 3 | < 1 | < 2 | 0.0611 | [15 empty] | 16
| 1 | 4 | < 1 | < 2 | 0.0527 | [15 empty] | 0.058217
| 1 | 5 | < 1 | < 2 | 0.0606 | [15 empty] | 8
| 1 | 6 | < 1 | < 2 | 0.0593
| Mean |  |  |  | 0.058217
| Median |  |  |  | 0.05865
| Std Dev. |  |  |  | 0.0030341
| Rel.Std.Dev. |  |  |  | 5.212%
| PDM3 |  |  |  | -5.551e-14%
| Z-Score (Absolute) |  | 0 | 0.76 | 0.67
| Analytical results for Ho in OREAS 254c (Certified Value 0.61 ppm) | [19 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.62 | 0.65284 | 0.70482 | 0.6 | 0.6 | 0.612 | 0.52 | [11 empty] | 1
| 1 | 2 | 0.59 | 0.62089 | 0.705 | 0.7 | 0.61 | 0.6176 | 0.56 | [11 empty] | 9
| 1 | 3 | 0.59 | 0.61508 | 0.74045 | 0.6 | 0.61 | 0.6049 | 0.54 | [11 empty] | 16
| 1 | 4 | 0.6 | 0.60322 | 0.62455 | 0.6 | 0.63 | 0.5958 | 0.53 | [11 empty] | 0.61068
| 1 | 5 | 0.59 | 0.60469 | 0.65394 | 0.6 | 0.63 | 0.6194 | 0.56 | [11 empty] | 35
| 1 | 6 | 0.58 | 0.62672 | 0.66582 | 0.6 | 0.6 | 0.6269 | 0.56
| Mean |  | 0.595 | 0.62057 | 0.68243 | 0.61667 | 0.61333 | 0.61277 | 0.545
| Median |  | 0.59 | 0.61799 | 0.68532 | 0.6 | 0.61 | 0.6148 | 0.55
| Std Dev. |  | 0.013784 | 0.018234 | 0.041988 | 0.040825 | 0.013663 | 0.011113 | 0.017607
| Rel.Std.Dev. |  | 2.317% | 2.938% | 6.153% | 6.62% | 2.228% | 1.814% | 3.231%
| PDM3 |  | -2.568% | 1.62% | 11.75% | 0.9797% | 0.4339% | 0.3411% | -10.76%
| Z-Score (Absolute) |  | 0.95 | 0.4 | 3.66 | NA | 0.01 | 0.01 | 3.59
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 254c (Certified Value 0.06 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.07 | 0.05 | < 0.1 | < 2 | 0.095 | 0.08 | 0.0668 | 0.061 | < 0.05 | < 0.1 | 0.058 | 0.057 | 0.055 | 0.061 |  |  |  |  | 1
| 1 | 2 | 0.05 | 0.05 | < 0.1 | < 2 | 0.058 | < 0.05 | 0.0688 | 0.057 | < 0.05 | < 0.1 | 0.055 | 0.056 | 0.054 | 0.066 |  |  |  |  | 22
| 1 | 3 | 0.06 | 0.05 | < 0.1 | < 2 | 0.048 | 0.06 | 0.0635 | 0.064 | 0.074177 | < 0.1 | 0.056 | 0.057 | 0.053 | 0.059 |  |  |  |  | 16
| 1 | 4 | 0.05 | 0.05 | < 0.1 | < 2 | 0.054 | 0.05 | 0.0652 | 0.049 | 0.072872 | < 0.1 | 0.057 | 0.056 | 0.055 | 0.062 |  |  |  |  | 0.059506
| 1 | 5 | 0.05 | 0.06 | < 0.1 | < 2 | 0.054 | < 0.05 | 0.0656 | 0.073 | 0.056758 | < 0.1 | 0.057 | 0.058 | 0.056 | 0.063 |  |  |  |  | 36
| 1 | 6 | 0.05 | 0.06 | < 0.1 | < 2 | 0.063 | 0.07 | 0.0686 | 0.071 | 0.064117 | < 0.1 | 0.057 | 0.055 | 0.06 | 0.062
| Mean |  | 0.055 | 0.053333 |  |  | 0.062 | 0.065 | 0.066417 | 0.0625 | 0.066981 |  | 0.056667 | 0.0565 | 0.0555 | 0.062167
| Median |  | 0.05 | 0.05 |  |  | 0.056 | 0.065 | 0.0662 | 0.0625 | 0.068495 |  | 0.057 | 0.0565 | 0.055 | 0.062
| Std Dev. |  | 0.0083666 | 0.005164 |  |  | 0.016912 | 0.01291 | 0.0020615 | 0.0089387 | 0.0081485 |  | 0.0010328 | 0.0010488 | 0.002429 | 0.0023166
| Rel.Std.Dev. |  | 15.21% | 9.682% |  |  | 27.28% | 19.86% | 3.104% | 14.3% | 12.17% |  | 1.823% | 1.856% | 4.377% | 3.726%
| PDM3 |  | -7.572% | -10.37% |  |  | 4.191% | 9.233% | 11.61% | 5.032% | 12.56% |  | -4.771% | -5.051% | -6.732% | 4.471%
| Z-Score (Absolute) |  | 0.13 | 0.35 | 0.78 | 123.2 | 0.78 | 0.57 | 1.36 | 0.85 | 0.39 | 0.78 | 0.09 | 0.07 | 0.07 | 0.8
| Analytical results for K in OREAS 254c (Certified Value 1.29 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.2844 | 1.3595 | 1.325 | 1.2076 | 1.27 | 1.3 | 1.36 | 1.26 | 1.28 | 1.29 | 1.2956 | 1.27 | 1.2615 | 1.28 | 1.38 | 1.31 | 1.22 | 1.37 | 1
| 1 | 2 | 1.288 | 1.3654 | 1.3029 | 1.1949 | 1.26 | 1.35 | 1.37 | 1.24 | 1.24 | 1.25 | 1.2917 | 1.27 | 1.2735 | 1.25 | 1.38 | 1.27 | 1.23 | 1.38 | #N/A
| 1 | 3 | 1.2681 | 1.3482 | 1.3021 | 1.2178 | 1.25 | 1.33 | 1.35 | 1.29 | 1.24 | 1.31 | 1.2639 | 1.25 | 1.2664 | 1.27 | 1.35 | 1.29 | 1.22 | 1.42 | 16
| 1 | 4 | 1.2696 | 1.3435 | 1.3076 | 1.194 | 1.25 | 1.38 | 1.36 | 1.29 | 1.24 | 1.28 | 1.2833 | 1.26 | 1.2675 | 1.31 | 1.32 | 0.96 | 1.25 | 1.36 | 1.2856
| 1 | 5 | 1.2539 | 1.3558 | 1.3286 | 1.1994 | 1.27 | 1.3 | 1.35 | 1.28 | 1.23 | 1.32 | 1.2687 | 1.31 | 1.2676 | 1.27 | 1.36 | 1.3 | 1.23 | 1.37 | 37
| 1 | 6 | 1.2644 | 1.365 | 1.3306 | 1.237 | 1.26 | 1.18 | 1.36 | 1.29 | 1.25 | 1.29 | 1.2862 | 1.27 | 1.2731 | 1.29 | 1.36 | 1.29 | 1.23 | 1.43
| Mean |  | 1.2714 | 1.3562 | 1.3161 | 1.2084 | 1.26 | 1.3067 | 1.3583 | 1.275 | 1.2467 | 1.29 | 1.2816 | 1.2717 | 1.2683 | 1.2783 | 1.3583 | 1.2367 | 1.23 | 1.3883
| Median |  | 1.2689 | 1.3577 | 1.3163 | 1.2035 | 1.26 | 1.315 | 1.36 | 1.285 | 1.24 | 1.29 | 1.2848 | 1.27 | 1.2675 | 1.275 | 1.36 | 1.29 | 1.23 | 1.375
| Std Dev. |  | 0.012759 | 0.0089387 | 0.013319 | 0.016604 | 0.0089443 | 0.069186 | 0.0075277 | 0.020736 | 0.017512 | 0.024495 | 0.012662 | 0.020412 | 0.004506 | 0.020412 | 0.022286 | 0.13619 | 0.010954 | 0.029269
| Rel.Std.Dev. |  | 1.004% | 0.6591% | 1.012% | 1.374% | 0.7099% | 5.295% | 0.5542% | 1.626% | 1.405% | 1.899% | 0.988% | 1.605% | 0.3553% | 1.597% | 1.641% | 11.01% | 0.8906% | 2.108%
| PDM3 |  | -1.102% | 5.497% | 2.376% | -5.999% | -1.989% | 1.641% | 5.66% | -0.8224% | -3.026% | 0.3444% | -0.3099% | -1.082% | -1.347% | -0.5631% | 5.66% | -3.804% | -4.323% | 7.993%
| Z-Score (Absolute) |  | 0.12 | 1.79 | 0.89 | 1.53 | 0.37 | 0.67 | 1.84 | 0.04 | 0.67 | 0.3 | 0.11 | 0.11 | 0.19 | 0.04 | 1.84 | 0.9 | 1.05 | 2.51
| Analytical results for La in OREAS 254c (Certified Value 37.9 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 37.19 | 38.453 | 38.9958 | 34.7 | 34.4 | 36.7731 | 38.8 | 38.6 | 30.6807 | 38.7 | 39 | 37.8 | 39.8 | 37.8 | 44.85 |  |  |  | 1
| 1 | 2 | 37 | 39.6693 | 39.6287 | 35.9 | 34.3 | 38.2103 | 38.5 | 37.8 | 31.07 | 37.8 | 38.3 | 37.9 | 39.3 | 37.3 | 47.47 |  |  |  | 23
| 1 | 3 | 37.11 | 39.3823 | 39.5044 | 35.2 | 35.3 | 37.0581 | 37.8 | 38.4 | 30.4022 | 39 | 39.3 | 38.2 | 39.6 | 36.2 | 47.09 |  |  |  | 16
| 1 | 4 | 38.03 | 38.2954 | 39.1059 | 39.2 | 36.6 | 37.6044 | 38.1 | 37.6 | 30.8792 | 36.5 | 39.9 | 36.7 | 38.2 | 39.2 | 45.6 |  |  |  | 37.8871
| 1 | 5 | 37.46 | 39.3822 | 39.9312 | 35.5 | 36.5 | 37.5536 | 37.4 | 37.7 | 30.2899 | 43.6 | 38.8 | 37.9 | 38.6 | 39 | 47.33 |  |  |  | 38
| 1 | 6 | 36.81 | 38.8235 | 39.4832 | 35.6 | 34.1 | 37.3767 | 40.6 | 36.3 | 31.0501 | 40.9 | 39.1 | 38.9 | 39.2 | 37.8 | 46.23
| Mean |  | 37.2667 | 39.0009 | 39.4415 | 36.0167 | 35.2 | 37.4294 | 38.5333 | 37.7333 | 30.7287 | 39.4167 | 39.0667 | 37.9 | 39.1167 | 37.8833 | 46.4283
| Median |  | 37.15 | 39.1028 | 39.4938 | 35.55 | 34.85 | 37.4651 | 38.3 | 37.75 | 30.7799 | 38.85 | 39.05 | 37.9 | 39.25 | 37.8 | 46.66
| Std Dev. |  | 0.43131 | 0.55982 | 0.34401 | 1.6117 | 1.1243 | 0.4954 | 1.1272 | 0.80911 | 0.32979 | 2.5103 | 0.53166 | 0.71274 | 0.608 | 1.1107 | 1.0538
| Rel.Std.Dev. |  | 1.157% | 1.435% | 0.8722% | 4.475% | 3.194% | 1.324% | 2.925% | 2.144% | 1.073% | 6.369% | 1.361% | 1.881% | 1.554% | 2.932% | 2.27%
| PDM3 |  | -1.637% | 2.94% | 4.103% | -4.937% | -7.092% | -1.208% | 1.706% | -0.4058% | -18.89% | 4.037% | 3.113% | 0.03414% | 3.245% | -0.009845% | 22.54%
| Z-Score (Absolute) |  | 0.37 | 0.64 | 0.89 | 1.09 | 1.56 | 0.27 | 0.37 | 0.1 | 4.14 | 0.88 | 0.67 | 0 | 0.7 | 0.01 | 4.93
| Analytical results for Li in OREAS 254c (Certified Value 24.1 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 24.6 | 26.0521 | 24.8389 | 40.9713 | 23 | 25 | 26.7 | 22.4878 | 24.7 | 23.9 | 21.294 | 24.6 | 24.8857 | 22.8 | 25.5 | 25.7 | 24.4 | 21 | 1
| 1 | 2 | 24.2 | 26.5322 | 24.3568 | 40.6859 | 23.5 | 25 | 25.9 | 22.6097 | 23.9 | 23.3 | 21.45 | 25.1 | 23.2046 | 22.8 | 25.5 | 25.5 | 24.5 | 21 | #N/A
| 1 | 3 | 23.9 | 26.3322 | 24.6922 | 39.9136 | 23.3 | 25 | 24.4 | 23.3121 | 23.9 | 24 | 20.982 | 24.8 | 25.4463 | 23.5 | 24.8 | 25.5 | 24.3 |  | 16
| 1 | 4 | 23.7 | 26.1958 | 24.7173 | 39.6651 | 22.7 | 25 | 25.4 | 23.5707 | 23.8 | 23.2 | 21.372 | 25.1 | 23.2455 | 23.9 | 24.5 | 20 | 24.9 | 21 | 24.1261
| 1 | 5 | 23.6 | 26.1903 | 24.7335 | 39.7213 | 23.2 | 26 | 25 | 23.1769 | 23.7 | 24.2 | 20.826 | 25.5 | 25.1694 | 23.4 | 25.3 | 25.7 | 24.5 | 22 | 39
| 1 | 6 | 24 | 26.3769 | 24.6489 | 40.8103 | 23.2 | 26 | 25.8 | 23.2189 | 24.2 | 24.5 | 21.45 | 25 | 23.7142 | 23.6 | 25.7 | 25.5 | 24.5 | 21
| Mean |  | 24 | 26.2799 | 24.6646 | 40.2946 | 23.15 | 25.3333 | 25.5333 | 23.0627 | 24.0333 | 23.85 | 21.229 | 25.0167 | 24.2776 | 23.3333 | 25.2167 | 24.65 | 24.5167 | 21.2
| Median |  | 23.95 | 26.264 | 24.7048 | 40.2998 | 23.2 | 25 | 25.6 | 23.1979 | 23.9 | 23.95 | 21.333 | 25.05 | 24.2999 | 23.45 | 25.4 | 25.5 | 24.5 | 21
| Std Dev. |  | 0.36332 | 0.169 | 0.1635 | 0.59112 | 0.27386 | 0.5164 | 0.79415 | 0.42271 | 0.36697 | 0.50892 | 0.26297 | 0.30605 | 1.0065 | 0.44572 | 0.46655 | 2.2801 | 0.20412 | 0.44721
| Rel.Std.Dev. |  | 1.514% | 0.6431% | 0.6629% | 1.467% | 1.183% | 2.038% | 3.11% | 1.833% | 1.527% | 2.134% | 1.239% | 1.223% | 4.146% | 1.91% | 1.85% | 9.25% | 0.8326% | 2.109%
| PDM3 |  | -0.5227% | 8.927% | 2.232% | 67.02% | -4.046% | 5.004% | 5.833% | -4.408% | -0.3845% | -1.144% | -12.01% | 3.691% | 0.628% | -3.286% | 4.52% | 2.171% | 1.619% | -12.13%
| Z-Score (Absolute) |  | 0.31 | 1.45 | 0.21 | 12.21 | 0.96 | 0.72 | 0.87 | 1.03 | 0.28 | 0.42 | 2.43 | 0.48 | 0.09 | 0.82 | 0.63 | 0.19 | 0.09 | 2.46
| Analytical results for Lu in OREAS 254c (Certified Value 0.17 ppm) | [19 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.18 | 0.17618 | 0.19931 | 0.2 | 0.18 | 0.06 | 0.14 | [11 empty] | 1
| 1 | 2 | 0.18 | 0.17243 | 0.19764 | 0.2 | 0.17 | 0.0671 | 0.14 | [11 empty] | 4
| 1 | 3 | 0.18 | 0.17566 | 0.18832 | 0.2 | 0.18 | 0.0578 | 0.15 | [11 empty] | 16
| 1 | 4 | 0.18 | 0.18265 | 0.18929 | 0.2 | 0.19 | < 0.05 | 0.14 | [11 empty] | 0.17475
| 1 | 5 | 0.18 | 0.18077 | 0.17879 | 0.2 | 0.19 | 0.0595 | 0.15 | [11 empty] | 40
| 1 | 6 | 0.17 | 0.17839 | 0.19319 | 0.2 | 0.18 | 0.0635 | 0.15
| Mean |  | 0.17833 | 0.17768 | 0.19109 | 0.2 | 0.18167 | 0.06158 | 0.145
| Median |  | 0.18 | 0.17728 | 0.19124 | 0.2 | 0.18 | 0.06 | 0.145
| Std Dev. |  | 0.0040825 | 0.0037047 | 0.007442 | 3.0405e-17 | 0.0075277 | 0.0037171 | 0.0054772
| Rel.Std.Dev. |  | 2.289% | 2.085% | 3.894% | 1.52e-14% | 4.144% | 6.036% | 3.777%
| PDM3 |  | 2.048% | 1.674% | 9.349% | 14.45% | 3.955% | -64.76% | -17.03%
| Z-Score (Absolute) |  | 0.03 | 0.03 | 1.05 | NA | 0.29 | 9.87 | 2.66
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 254c (Certified Value 1.51 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.5129 | 1.6235 | 1.5461 | 1.6851 | 1.53 | 1.47 | 1.48 | 1.48 | 1.54 | 1.56 | 1.5136 | 1.48 | 1.4941 | 1.53 | 1.52 | 1.53 | 1.42 | 1.478 | 1
| 1 | 2 | 1.5144 | 1.6442 | 1.5253 | 1.6819 | 1.55 | 1.53 | 1.43 | 1.49 | 1.5 | 1.54 | 1.5195 | 1.52 | 1.5161 | 1.5 | 1.52 | 1.51 | 1.42 | 1.478 | #N/A
| 1 | 3 | 1.5001 | 1.6614 | 1.5348 | 1.6762 | 1.55 | 1.54 | 1.46 | 1.52 | 1.48 | 1.54 | 1.5008 | 1.49 | 1.5166 | 1.51 | 1.47 | 1.51 | 1.42 | 1.52 | 16
| 1 | 4 | 1.4993 | 1.6354 | 1.5224 | 1.6549 | 1.56 | 1.51 | 1.51 | 1.5 | 1.46 | 1.54 | 1.5177 | 1.51 | 1.4975 | 1.56 | 1.46 | 1.22 | 1.45 | 1.441 | 1.5093
| 1 | 5 | 1.4874 | 1.6588 | 1.552 | 1.6651 | 1.54 | 1.49 | 1.49 | 1.5 | 1.48 | 1.56 | 1.4965 | 1.54 | 1.5066 | 1.51 | 1.49 | 1.51 | 1.42 | 1.462 | 41
| 1 | 6 | 1.489 | 1.6635 | 1.5484 | 1.7079 | 1.57 | 1.47 | 1.47 | 1.51 | 1.5 | 1.5 | 1.5229 | 1.51 | 1.5082 | 1.52 | 1.52 | 1.51 | 1.41 | 1.521
| Mean |  | 1.5005 | 1.6478 | 1.5382 | 1.6785 | 1.55 | 1.5017 | 1.4733 | 1.5 | 1.4933 | 1.54 | 1.5119 | 1.5083 | 1.5065 | 1.5217 | 1.4967 | 1.465 | 1.4233 | 1.4833
| Median |  | 1.4997 | 1.6515 | 1.5405 | 1.6791 | 1.55 | 1.5 | 1.475 | 1.5 | 1.49 | 1.54 | 1.5157 | 1.51 | 1.5074 | 1.515 | 1.505 | 1.51 | 1.42 | 1.478
| Std Dev. |  | 0.011423 | 0.016184 | 0.01252 | 0.018205 | 0.014142 | 0.029944 | 0.027325 | 0.014142 | 0.027325 | 0.021909 | 0.010733 | 0.02137 | 0.0092964 | 0.02137 | 0.027325 | 0.12029 | 0.013663 | 0.031835
| Rel.Std.Dev. |  | 0.7613% | 0.9821% | 0.814% | 1.085% | 0.9124% | 1.994% | 1.855% | 0.9428% | 1.83% | 1.423% | 0.7099% | 1.417% | 0.6171% | 1.404% | 1.826% | 8.211% | 0.9599% | 2.146%
| PDM3 |  | -0.5815% | 9.177% | 1.912% | 11.21% | 2.697% | -0.5053% | -2.383% | -0.6157% | -1.057% | 2.035% | 0.17% | -0.06357% | -0.1834% | 0.8198% | -0.8366% | -2.935% | -5.695% | -1.72%
| Z-Score (Absolute) |  | 0.13 | 5.06 | 1.2 | 6.14 | 1.61 | 0.09 | 1.08 | 0.14 | 0.38 | 1.26 | 0.27 | 0.15 | 0.09 | 0.62 | 0.26 | 1.38 | 2.84 | 0.73
| Analytical results for Mn in OREAS 254c (Certified Value 0.046 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0472 | 0.045746 | 0.047074 | < 0.1 | 0.0445 | 0.0429 | 0.0466 | 0.0455 | 0.0472 | 0.0462 | 0.040871 | 0.0455 | 0.045411 | 0.0463 | 0.047 | 0.0483 | 0.0444 | 0.0442 | 1
| 1 | 2 | 0.0473 | 0.045221 | 0.046069 | < 0.1 | 0.0452 | 0.044 | 0.0467 | 0.0453 | 0.0458 | 0.0461 | 0.041195 | 0.047 | 0.046499 | 0.0457 | 0.047 | 0.0475 | 0.0448 | 0.0444 | 25
| 1 | 3 | 0.0468 | 0.04477 | 0.046709 | < 0.1 | 0.0449 | 0.0441 | 0.0468 | 0.0462 | 0.0453 | 0.0462 | 0.040764 | 0.0459 | 0.046445 | 0.0469 | 0.0456 | 0.0477 | 0.0451 | 0.046 | 16
| 1 | 4 | 0.0467 | 0.044581 | 0.046379 | < 0.1 | 0.0447 | 0.0488 | 0.0473 | 0.0461 | 0.0453 | 0.0456 | 0.040978 | 0.047 | 0.04535 | 0.0451 | 0.0451 | 0.0379 | 0.0453 | 0.0442 | 0.045808
| 1 | 5 | 0.0462 | 0.044981 | 0.046586 | < 0.1 | 0.044 | 0.0427 | 0.0469 | 0.046 | 0.0448 | 0.0467 | 0.04055 | 0.0473 | 0.046154 | 0.0462 | 0.0463 | 0.0477 | 0.0451 | 0.0442 | 42
| 1 | 6 | 0.0462 | 0.045215 | 0.047254 | < 0.1 | 0.0451 | 0.0434 | 0.0467 | 0.0453 | 0.0453 | 0.0445 | 0.041197 | 0.0463 | 0.04568 | 0.0462 | 0.0468 | 0.0475 | 0.0445 | 0.0461
| Mean |  | 0.046733 | 0.045086 | 0.046679 |  | 0.044733 | 0.044317 | 0.046833 | 0.045733 | 0.045617 | 0.045883 | 0.040926 | 0.0465 | 0.045923 | 0.046067 | 0.0463 | 0.0461 | 0.044867 | 0.04485
| Median |  | 0.04675 | 0.045098 | 0.046648 |  | 0.0448 | 0.0437 | 0.04675 | 0.04575 | 0.0453 | 0.04615 | 0.040925 | 0.04665 | 0.045917 | 0.0462 | 0.04655 | 0.0476 | 0.04495 | 0.0443
| Std Dev. |  | 0.00047188 | 0.00040872 | 0.00043776 |  | 0.00044121 | 0.0022675 | 0.00025033 | 0.00041312 | 0.00083766 | 0.00076267 | 0.0002526 | 0.00071274 | 0.00051128 | 0.00060882 | 0.00079498 | 0.0040279 | 0.00036148 | 0.00093327
| Rel.Std.Dev. |  | 1.01% | 0.9065% | 0.9378% |  | 0.9863% | 5.117% | 0.5345% | 0.9033% | 1.836% | 1.662% | 0.6172% | 1.533% | 1.113% | 1.322% | 1.717% | 8.737% | 0.8057% | 2.081%
| PDM3 |  | 2.021% | -1.576% | 1.901% |  | -2.346% | -3.255% | 2.239% | -0.1625% | -0.4172% | 0.165% | -10.66% | 1.511% | 0.2524% | 0.5652% | 1.075% | 0.638% | -2.054% | -2.091%
| Z-Score (Absolute) |  | 0.7 | 0.69 | 0.66 | 3.47 | 0.99 | 1.34 | 0.79 | 0.14 | 0.24 | 0.02 | 4.21 | 0.51 | 0.02 | 0.14 | 0.34 | 0.17 | 0.88 | 0.89
| Analytical results for Mo in OREAS 254c (Certified Value 3.87 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 3.8 | 3.9653 | 4.9822 | 3.9 | < 1 | 4.27 | 3.8688 | 3.99 | 3.9 | 4.184 | 3.53 | 3.3333 | 3.61 | 3.63 | 4.05 | 3.9 | 3.71 |  | 1
| 1 | 2 | 4 | 3.9196 | 5.1016 | 3.9 | < 1 | 4.29 | 3.8795 | 3.73 | 3.7 | 4.0673 | 3.54 | 3.6927 | 3.81 | 3.61 | 3.94 | 3.93 | 3.86 |  | 26
| 1 | 3 | 3.8 | 3.8627 | 4.6432 | 3.9 | 4 | 4.33 | 4.2557 | 3.83 | 3.7 | 4.154 | 3.94 | 3.6592 | 3.66 | 3.65 | 4.03 | 3.74 | 4.09 |  | 16
| 1 | 4 | 3.7 | 3.7398 | 5.0909 | 4 | < 1 | 4.34 | 4.1682 | 3.94 | 3.6 | 3.9888 | 3.44 | 3.3014 | 3.6 | 3.46 | 3.92 | 4.14 | 3.8 |  | 3.8664
| 1 | 5 | 3.8 | 3.9264 | 5.253 | 3.9 | 1 | 4.33 | 3.8767 | 3.68 | 3.9 | 4.2207 | 3.82 | 3.6062 | 3.72 | 3.59 | 3.92 | 3.99 | 3.95 |  | 43
| 1 | 6 | 3.6 | 3.9214 | 5.1331 | 3.9 | 2 | 4.3 | 4.0472 | 3.99 | 3.8 | 4.1807 | 3.64 | 3.4555 | 3.76 | 3.69 | 4 | 3.78 | 4.08
| Mean |  | 3.7833 | 3.8892 | 5.034 | 3.9167 | 2.3333 | 4.31 | 4.016 | 3.86 | 3.7667 | 4.1326 | 3.6517 | 3.508 | 3.6933 | 3.605 | 3.9767 | 3.9133 | 3.915
| Median |  | 3.8 | 3.9205 | 5.0963 | 3.9 | 2 | 4.315 | 3.9633 | 3.885 | 3.75 | 4.1673 | 3.59 | 3.5308 | 3.69 | 3.62 | 3.97 | 3.915 | 3.905
| Std Dev. |  | 0.13292 | 0.080186 | 0.21024 | 0.040825 | 1.5275 | 0.027568 | 0.16809 | 0.13446 | 0.12111 | 0.087313 | 0.19167 | 0.16884 | 0.084301 | 0.07893 | 0.057504 | 0.14528 | 0.1532
| Rel.Std.Dev. |  | 3.513% | 2.062% | 4.176% | 1.042% | 65.47% | 0.6396% | 4.185% | 3.483% | 3.215% | 2.113% | 5.249% | 4.813% | 2.283% | 2.189% | 1.446% | 3.712% | 3.913%
| PDM3 |  | -2.149% | 0.5895% | 30.2% | 1.299% | -39.65% | 11.47% | 3.869% | -0.1662% | -2.58% | 6.884% | -5.554% | -9.269% | -4.477% | -6.761% | 2.851% | 1.213% | 1.256%
| Z-Score (Absolute) |  | 0.56 | 0 | 6.09 | 0.15 | 13.15 | 2.24 | 0.67 | 0.16 | 0.65 | 1.29 | 1.26 | 2.03 | 1.04 | 1.51 | 0.47 | 0.13 | 0.14
| Analytical results for Na in OREAS 254c (Certified Value 0.969 wt.%) | [19 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.9603 | 0.85507 | 0.98566 | 2.62 | 0.953 | 0.95 | 0.98 | 0.99 | 0.98 | 0.992 | 0.9587 | 0.92 | 0.95203 | 0.99 | 0.97 | 0.99 | 0.93 | 1.006 | 1
| 1 | 2 | 0.965 | 0.8583 | 0.98172 | 2.5663 | 0.948 | 0.97 | 0.97 | 0.98 | 0.95 | 0.966 | 0.9577 | 0.96 | 0.97271 | 0.97 | 0.98 | 0.98 | 0.93 | 1.018 | #N/A
| 1 | 3 | 0.9538 | 0.85604 | 0.97418 | 2.4653 | 0.956 | 0.97 | 0.98 | 1.01 | 0.94 | 1.001 | 0.9448 | 0.94 | 0.95374 | 0.97 | 0.94 | 0.98 | 0.95 | 1.049 | 16
| 1 | 4 | 0.9478 | 0.85754 | 0.98606 | 2.4611 | 0.948 | 0.98 | 1.01 | 1 | 0.93 | 0.973 | 0.9562 | 0.95 | 0.99129 | 0.99 | 0.93 | 0.74 | 0.95 | 0.995 | 0.9694
| 1 | 5 | 0.9517 | 0.85154 | 0.9835 | 2.4852 | 0.944 | 0.97 | 1 | 0.99 | 0.94 | 0.995 | 0.944 | 0.97 | 0.99359 | 0.98 | 0.96 | 0.98 | 0.93 | 1 | 44
| 1 | 6 | 0.947 | 0.86238 | 0.99175 | 2.5431 | 0.948 | 0.94 | 0.98 | 1.01 | 0.95 | 0.973 | 0.9567 | 0.96 | 0.9813 | 0.97 | 0.97 | 0.98 | 0.93 | 1.041
| Mean |  | 0.95427 | 0.85681 | 0.98381 | 2.5235 | 0.9495 | 0.96333 | 0.98667 | 0.99667 | 0.94833 | 0.98333 | 0.95302 | 0.95 | 0.97411 | 0.97833 | 0.95833 | 0.94167 | 0.93667 | 1.0182
| Median |  | 0.95275 | 0.85679 | 0.98458 | 2.5141 | 0.948 | 0.97 | 0.98 | 0.995 | 0.945 | 0.9825 | 0.95645 | 0.955 | 0.97701 | 0.975 | 0.965 | 0.98 | 0.93 | 1.012
| Std Dev. |  | 0.0071138 | 0.0036112 | 0.0058113 | 0.063707 | 0.0042778 | 0.015055 | 0.015055 | 0.012111 | 0.017224 | 0.014404 | 0.0067342 | 0.017889 | 0.018058 | 0.0098319 | 0.019408 | 0.098877 | 0.010328 | 0.022302
| Rel.Std.Dev. |  | 0.7455% | 0.4215% | 0.5907% | 2.525% | 0.4505% | 1.563% | 1.526% | 1.215% | 1.816% | 1.465% | 0.7066% | 1.883% | 1.854% | 1.005% | 2.025% | 10.5% | 1.103% | 2.19%
| PDM3 |  | -1.561% | -11.61% | 1.487% | 160.3% | -2.053% | -0.6256% | 1.781% | 2.813% | -2.173% | 1.437% | -1.69% | -2.001% | 0.4861% | 0.9217% | -1.141% | -2.861% | -3.376% | 5.031%
| Z-Score (Absolute) |  | 0.24 | 3.83 | 0.85 | 57.47 | 0.42 | 0.09 | 0.95 | 1.32 | 0.46 | 0.83 | 0.29 | 0.4 | 0.49 | 0.64 | 0.09 | 0.7 | 0.89 | 2.11
| Analytical results for Nb in OREAS 254c (Certified Value 30.5 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 30.27 | 34.7084 | 80.6783 | 29.8 | 0.6 | 23.9 | 28.6944 | 30.8 | 32.7 | 45.7674 | 30.1 | 35.2118 | 30 | 32.4 | 33 | 30 | 29 |  | 1
| 1 | 2 | 30.15 | 34.1024 | 58.4709 | 30 | 0.6 | 24.2 | 27.9867 | 29.8 | 31.5 | 46.1405 | 28.7 | 34.4701 | 29.2 | 31.6 | 31.9 | 30.4 | 30.7 |  | 5
| 1 | 3 | 29.75 | 32.9216 | 59.5557 | 30 | 0.4 | 25.3 | 28.7846 | 29.5 | 32.5 | 44.8955 | 29.3 | 34.4265 | 29.8 | 32 | 32.9 | 29.5 | 30.3 |  | 16
| 1 | 4 | 29.77 | 33.0971 | 57.855 | 30.3 | < 0.1 | 25.5 | 29.8893 | 29.4 | 31.3 | 43.8697 | 28.1 | 32.2889 | 30.3 | 30.8 | 32.2 | 32.3 | 29.8 |  | 30.5205
| 1 | 5 | 30.25 | 32.958 | 59.3527 | 29.9 | 0.6 | 25.5 | 28.8333 | 28.5 | 32.6 | 46.2591 | 33.4 | 34.1192 | 29.7 | 31.4 | 31.7 | 31.5 | 29.2 |  | 45
| 1 | 6 | 30.36 | 31.5724 | 62.7443 | 30.1 | 0.2 | 24.3 | 29.1536 | 29.5 | 33.1 | 46.8618 | 30.7 | 32.5503 | 30.6 | 32.5 | 32.6 | 30.3 | 29.5
| Mean |  | 30.0917 | 33.2267 | 63.1095 | 30.0167 | 0.48 | 24.7833 | 28.8903 | 29.5833 | 32.2833 | 45.6323 | 30.05 | 33.8445 | 29.9333 | 31.7833 | 32.3833 | 30.6667 | 29.75
| Median |  | 30.2 | 33.0276 | 59.4542 | 30 | 0.6 | 24.8 | 28.8089 | 29.5 | 32.55 | 45.954 | 29.7 | 34.2728 | 29.9 | 31.8 | 32.4 | 30.35 | 29.65
| Std Dev. |  | 0.26551 | 1.0845 | 8.7712 | 0.17224 | 0.17889 | 0.72778 | 0.62233 | 0.7414 | 0.71671 | 1.0799 | 1.8887 | 1.1636 | 0.48854 | 0.64627 | 0.53448 | 1.0367 | 0.65345
| Rel.Std.Dev. |  | 0.8823% | 3.264% | 13.9% | 0.5738% | 37.27% | 2.937% | 2.154% | 2.506% | 2.22% | 2.366% | 6.285% | 3.438% | 1.632% | 2.033% | 1.65% | 3.38% | 2.196%
| PDM3 |  | -1.405% | 8.867% | 106.8% | -1.651% | -98.43% | -18.8% | -5.341% | -3.07% | 5.776% | 49.51% | -1.541% | 10.89% | -1.924% | 4.138% | 6.104% | 0.479% | -2.524%
| Z-Score (Absolute) |  | 0 | 1.25 | 13.16 | 0.03 | 11.83 | 2.12 | 0.48 | 0.2 | 0.87 | 6.19 | 0.02 | 1.5 | 0.06 | 0.67 | 0.91 | 0.23 | 0.14
| Analytical results for Nd in OREAS 254c (Certified Value 32 ppm) | [19 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.14 | 1.22 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 1
1
| 1 | 1 | 32.31 | 33.3415 | 31 | 31.1 | 24.7878 | 39.7 | [12 empty] | 1
| 1 | 2 | 30.83 | 32.7276 | 31.9 | 31.3 | 25.18 | 39.9 | [12 empty] | 6
| 1 | 3 | 31.12 | 32.8187 | 31.4 | 31.6 | 24.6039 | 39.3 | [12 empty] | 16
| 1 | 4 | 31.62 | 32.0355 | 32.9 | 31.8 | 24.9969 | 38.7 | [12 empty] | 31.9647
| 1 | 5 | 30.97 | 32.9846 | 31.6 | 32.1 | 24.6316 | 39.4 | [12 empty] | 46
| 1 | 6 | 30.74 | 32.1895 | 31.3 | 31.4 | 25.2426 | 41.2
| Mean |  | 31.265 | 32.6829 | 31.6833 | 31.55 | 24.9071 | 39.7
| Median |  | 31.045 | 32.7731 | 31.5 | 31.5 | 24.8924 | 39.55
| Std Dev. |  | 0.59849 | 0.49148 | 0.66758 | 0.36194 | 0.27463 | 0.84143
| Rel.Std.Dev. |  | 1.914% | 1.504% | 2.107% | 1.147% | 1.103% | 2.119%
| PDM3 |  | -2.189% | 2.247% | -0.8803% | -1.297% | -22.08% | 24.2%
| Z-Score (Absolute) |  | 0.33 | 1.01 | 0.06 | 0.06 | 6.38 | 7.69
| Analytical results for Ni in OREAS 254c (Certified Value 86 ppm) | [19 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
1
| 1 | 1 | 88.1 | 88.8061 | 72.7208 | 82 | 89 | 88.9 | 80.1167 | 86.9 | 84.1 | 70.3533 | 79.5 | 85.973 | 88.1 | 86 | 89.5 | 84 | 86 |  | 1
| 1 | 2 | 85.9 | 87.1259 | 72.1862 | 83 | 90 | 91.7 | 79.74 | 84.1 | 83.9 | 70.8415 | 78.6 | 87.8476 | 87.7 | 85 | 88.5 | 84.1 | 88 |  | 29
| 1 | 3 | 86.5 | 87.8289 | 74.2784 | 82 | 89 | 88.7 | 82.67 | 83.2 | 83.9 | 69.5691 | 80.1 | 87.2471 | 88.5 | 86 | 89.2 | 89.6 | 91 |  | 16
| 1 | 4 | 86.5 | 86.8938 | 72.952 | 82 | 90 | 92.2 | 81.77 | 83.1 | 84.1 | 70.4558 | 77.5 | 83.1471 | 87.1 | 81.7 | 86.1 | 86.9 | 86 |  | 85.8479
| 1 | 5 | 84.1 | 88.1212 | 73.1682 | 82 | 90 | 92.9 | 80.9867 | 83.4 | 85.3 | 69.7798 | 86 | 86.2759 | 88.3 | 86.5 | 88.2 | 85.4 | 86 |  | 47
| 1 | 6 | 85.4 | 88.4302 | 73.8039 | 82 | 90 | 89.1 | 81.8367 | 84.5 | 81.4 | 70.5284 | 83 | 83.8926 | 86.8 | 86.4 | 88.7 | 83.8 | 91
| Mean |  | 86.0833 | 87.8677 | 73.1849 | 82.1667 | 89.6667 | 90.5833 | 81.1867 | 84.2 | 83.7833 | 70.2547 | 80.7833 | 85.7305 | 87.75 | 85.2667 | 88.3667 | 85.6333 | 88
| Median |  | 86.2 | 87.9751 | 73.0601 | 82 | 90 | 90.4 | 81.3783 | 83.75 | 84 | 70.4045 | 79.8 | 86.1244 | 87.9 | 86 | 88.6 | 84.75 | 87
| Std Dev. |  | 1.3303 | 0.74317 | 0.75448 | 0.40825 | 0.5164 | 1.8872 | 1.1172 | 1.4311 | 1.2813 | 0.4827 | 3.1556 | 1.8542 | 0.68044 | 1.8261 | 1.2061 | 2.2704 | 2.4495
| Rel.Std.Dev. |  | 1.545% | 0.8458% | 1.031% | 0.4969% | 0.5759% | 2.083% | 1.376% | 1.7% | 1.529% | 0.6871% | 3.906% | 2.163% | 0.7754% | 2.142% | 1.365% | 2.651% | 2.784%
| PDM3 |  | 0.2743% | 2.353% | -14.75% | -4.288% | 4.448% | 5.516% | -5.43% | -1.92% | -2.405% | -18.16% | -5.899% | -0.1367% | 2.216% | -0.677% | 2.934% | -0.2499% | 2.507%
| Z-Score (Absolute) |  | 0.13 | 0.64 | 3.55 | 0.99 | 1.15 | 1.41 | 1.27 | 0.41 | 0.53 | 4.38 | 1.38 | 0.03 | 0.6 | 0.1 | 0.78 | 0 | 0.67
| Analytical results for P in OREAS 254c (Certified Value 0.089 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0867 | 0.090879 | < 0.1 | 0.09 | 0.071 | 0.092 | 0.0869 | 0.09 | 0.089 | 0.082286 | 0.087 | 0.089286 | 0.09 | 0.09 | 0.091 | 0.084 | 0.086 |  | 1
| 1 | 2 | 0.087 | 0.088641 | < 0.1 | 0.09 | 0.072 | 0.0944 | 0.0868 | 0.087 | 0.089 | 0.082685 | 0.089 | 0.089687 | 0.09 | 0.09 | 0.09 | 0.084 | 0.086 |  | #N/A
| 1 | 3 | 0.086 | 0.090581 | < 0.1 | 0.09 | 0.069 | 0.093 | 0.0868 | 0.086 | 0.089 | 0.081987 | 0.089 | 0.090192 | 0.09 | 0.087 | 0.09 | 0.086 | 0.089 |  | 16
| 1 | 4 | 0.0853 | 0.087833 | < 0.1 | 0.09 | 0.069 | 0.0937 | 0.0868 | 0.086 | 0.089 | 0.082085 | 0.088 | 0.088749 | 0.09 | 0.086 | 0.073 | 0.085 | 0.086 |  | 0.088594
| 1 | 5 | 0.0842 | 0.089435 | < 0.1 | 0.09 | 0.07 | 0.0933 | 0.0864 | 0.087 | 0.091 | 0.081488 | 0.09 | 0.089444 | 0.09 | 0.088 | 0.09 | 0.084 | 0.087 |  | 48
| 1 | 6 | 0.086 | 0.090244 | < 0.1 | 0.091 | 0.069 | 0.0924 | 0.0871 | 0.088 | 0.087 | 0.082785 | 0.089 | 0.089719 | 0.09 | 0.089 | 0.091 | 0.083 | 0.089
| Mean |  | 0.085867 | 0.089602 |  | 0.090167 | 0.07 | 0.093133 | 0.0868 | 0.087333 | 0.089 | 0.082219 | 0.088667 | 0.089513 | 0.09 | 0.088333 | 0.0875 | 0.084333 | 0.087167
| Median |  | 0.086 | 0.08984 |  | 0.09 | 0.0695 | 0.09315 | 0.0868 | 0.087 | 0.089 | 0.082186 | 0.089 | 0.089565 | 0.09 | 0.0885 | 0.09 | 0.084 | 0.0865
| Std Dev. |  | 0.0010113 | 0.0011901 |  | 0.00040825 | 0.0012649 | 0.00087101 | 0.00022804 | 0.0015055 | 0.0012649 | 0.00047921 | 0.0010328 | 0.00048454 | 1.5202e-17 | 0.001633 | 0.0071204 | 0.0010328 | 0.001472
| Rel.Std.Dev. |  | 1.178% | 1.328% |  | 0.4528% | 1.807% | 0.9352% | 0.2627% | 1.724% | 1.421% | 0.5828% | 1.165% | 0.5413% | 1.689e-14% | 1.849% | 8.138% | 1.225% | 1.689%
| PDM3 |  | -3.078% | 1.138% |  | 1.775% | -20.99% | 5.124% | -2.025% | -1.423% | 0.4584% | -7.195% | 0.0821% | 1.037% | 1.587% | -0.2941% | -1.235% | -4.809% | -1.611%
| Z-Score (Absolute) |  | 0.55 | 0.7 | 12.56 | 0.89 | 5.86 | 1.89 | 0.23 | 0.06 | 0.5 | 1.77 | 0.39 | 0.67 | 0.84 | 0.28 | 0 | 1.06 | 0.11
| Analytical results for Pb in OREAS 254c (Certified Value 10.7 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 11.5 | 11.3544 | 5.8522 | 12 | 9 | 8.7 | 9.3927 | 11 | 11 | 8.5622 | 11.1 | 10.2226 | 11.1 | 10.4 | 11.4 | 10.8 | 13.1 |  | 1
| 1 | 2 | 11.2 | 12.0052 | 5.9444 | 13 | 10 | 8.9 | 9.8788 | 10.8 | 10.8 | 8.6735 | 11 | 10.431 | 11.3 | 10.3 | 11.2 | 11 | 13.5 |  | 30
| 1 | 3 | 11.1 | 12.049 | 6.7049 | 12 | 11 | 9.1 | 9.6584 | 10.8 | 11.1 | 8.0878 | 12.7 | 10.6552 | 11 | 10.1 | 11.4 | 9.8 | 14.3 |  | 16
| 1 | 4 | 11.3 | 12.1386 | 4.2218 | 12 | 10 | 8.5 | 11.2329 | 10.7 | 11.5 | 8.2753 | 10.6 | 9.0231 | 11.5 | 10 | 11.2 | 11.5 | 15.4 |  | 10.7207
| 1 | 5 | 11.4 | 12.9958 | 3.3812 | 12 | 10 | 8.8 | 10.2 | 10.7 | 11.2 | 8.6993 | 11.3 | 10.4137 | 11.2 | 10.1 | 11.2 | 11.3 | 14.6 |  | 49
| 1 | 6 | 11 | 11.5043 | 6.2241 | 12 | 11 | 8.5 | 10.2901 | 10.4 | 11 | 8.174 | 11.5 | 9.045 | 11.1 | 10.7 | 11.4 | 10.8 | 13
| Mean |  | 11.25 | 12.0079 | 5.3881 | 12.1667 | 10.1667 | 8.75 | 10.1088 | 10.7333 | 11.1 | 8.412 | 11.3667 | 9.9651 | 11.2 | 10.2667 | 11.3 | 10.8667 | 13.9833
| Median |  | 11.25 | 12.0271 | 5.8983 | 12 | 10 | 8.75 | 10.0394 | 10.75 | 11.05 | 8.4187 | 11.2 | 10.3181 | 11.15 | 10.2 | 11.3 | 10.9 | 13.9
| Std Dev. |  | 0.18708 | 0.57884 | 1.2919 | 0.40825 | 0.75277 | 0.23452 | 0.64395 | 0.19664 | 0.23664 | 0.26604 | 0.72019 | 0.73413 | 0.17889 | 0.2582 | 0.10954 | 0.59217 | 0.94534
| Rel.Std.Dev. |  | 1.663% | 4.821% | 23.98% | 3.355% | 7.404% | 2.68% | 6.37% | 1.832% | 2.132% | 3.163% | 6.336% | 7.367% | 1.597% | 2.515% | 0.9694% | 5.449% | 6.76%
| PDM3 |  | 4.937% | 12.01% | -49.74% | 13.49% | -5.168% | -18.38% | -5.707% | 0.1178% | 3.538% | -21.53% | 6.025% | -7.048% | 4.471% | -4.235% | 5.404% | 1.361% | 30.43%
| Z-Score (Absolute) |  | 0.43 | 1.28 | 6.16 | NA | NA | 2.38 | 0.85 | 0.15 | 0.26 | 2.76 | 0.56 | 1.01 | 0.37 | 0.67 | 0.49 | 0 | 3.5
Results from laboratories 1.07 and 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 254c (Certified Value 8.49 ppm) | [19 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 8.34 | 8.738 | 9.7618 | 8 | 8.06 | 4.4371 | 9.83 | [11 empty] | 1
| 1 | 2 | 8.3 | 8.6844 | 9.2203 | 8.4 | 7.94 | 4.5667 | 9.85 | [11 empty] | 8
| 1 | 3 | 8.32 | 8.8816 | 9.1191 | 8.1 | 8.14 | 4.3845 | 9.83 | [11 empty] | 16
| 1 | 4 | 8.3 | 8.4738 | 9.0695 | 8.4 | 8.37 | 4.4564 | 9.53 | [11 empty] | 8.491
| 1 | 5 | 8.36 | 8.6836 | 9.3018 | 8 | 8.34 | 4.3757 | 9.94 | [11 empty] | 50
| 1 | 6 | 8.18 | 8.7694 | 9.0307 | 8.2 | 7.86 | 4.3324 | 10.18
| Mean |  | 8.3 | 8.7051 | 9.2505 | 8.1833 | 8.1183 | 4.4255 | 9.86
| Median |  | 8.31 | 8.7112 | 9.1697 | 8.15 | 8.1 | 4.4108 | 9.84
| Std Dev. |  | 0.063246 | 0.13473 | 0.26955 | 0.18348 | 0.20731 | 0.082276 | 0.20957
| Rel.Std.Dev. |  | 0.762% | 1.548% | 2.914% | 2.242% | 2.554% | 1.859% | 2.125%
| PDM3 |  | -2.25% | 2.522% | 8.945% | -3.624% | -4.389% | -47.88% | 16.12%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.58 | 0.19 | 0.3 | 6.45 | 2.6
| Analytical results for Rb in OREAS 254c (Certified Value 65 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 65.68 | 69.0041 | 62.8 | 57.8 | 66.7 | 66.9 | 61 | 42.5754 | 62 | 63.1571 | 60.4 | 65.7 | 68.6 | 63.4 | 73.4 |  |  |  | 1
| 1 | 2 | 65.01 | 68.6221 | 63.1 | 59.4 | 68.6 | 64.8 | 57.6 | 44.6249 | 59.4 | 64.3114 | 59 | 67.1 | 65.7 | 63.9 | 75.7 |  |  |  | 32
| 1 | 3 | 63.98 | 69.3357 | 63.3 | 58.7 | 68.8 | 64.1 | 61.1 | 40.4955 | 59.1 | 64.1143 | 60.5 | 65.9 | 68.3 | 61.9 | 74.3 |  |  |  | 16
| 1 | 4 | 64.42 | 68.0058 | 63.4 | 65.1 | 69.9 | 64 | 58.5 | 42.5317 | 57.7 | 60.7871 | 62.1 | 64.1 | 66.5 | 66.5 | 74 |  |  |  | 64.5783
| 1 | 5 | 64.5 | 69.5936 | 63.1 | 58 | 69.6 | 61.2 | 60.7 | 41.8966 | 65.5 | 62.8507 | 59.8 | 65.7 | 67.2 | 65.4 | 77.5 |  |  |  | 51
| 1 | 6 | 63.93 | 69.5198 | 64.3 | 55.9 | 66 | 64.6 | 62.2 | 41.8003 | 63 | 59.8966 | 61.4 | 68 | 68.3 | 63.6 | 73.2
| Mean |  | 64.5867 | 69.0135 | 63.3333 | 59.15 | 68.2667 | 64.2667 | 60.1833 | 42.3207 | 61.1167 | 62.5195 | 60.5333 | 66.0833 | 67.4333 | 64.1167 | 74.6833
| Median |  | 64.46 | 69.1699 | 63.2 | 58.35 | 68.7 | 64.35 | 60.85 | 42.2141 | 60.7 | 63.0039 | 60.45 | 65.8 | 67.75 | 63.75 | 74.15
| Std Dev. |  | 0.6647 | 0.61081 | 0.5164 | 3.1425 | 1.5769 | 1.8348 | 1.752 | 1.3565 | 2.9061 | 1.7973 | 1.1057 | 1.3393 | 1.1656 | 1.6167 | 1.6388
| Rel.Std.Dev. |  | 1.029% | 0.8851% | 0.8154% | 5.313% | 2.31% | 2.855% | 2.911% | 3.205% | 4.755% | 2.875% | 1.827% | 2.027% | 1.729% | 2.521% | 2.194%
| PDM3 |  | 0.01294% | 6.868% | -1.928% | -8.406% | 5.711% | -0.4826% | -6.806% | -34.47% | -5.36% | -3.188% | -6.264% | 2.331% | 4.421% | -0.7149% | 15.65%
| Z-Score (Absolute) |  | 0.1 | 1 | 0.16 | 1.01 | 0.84 | 0.03 | 0.8 | 4.43 | 0.61 | 0.32 | 0.73 | 0.4 | 0.67 | 0 | 2.15
| Analytical results for Re in OREAS 254c (Certified Value < 0.002 ppm) | [19 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | 0.004 | [5 empty] | 1
| 1 | 2 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | 0.003 | [5 empty] | 33
| 1 | 3 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | 0.002 | [5 empty] | 16
| 1 | 4 | < 0.002 | < 0.05 | 0.007 | < 0.002 | 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | 0.003 | [5 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | 0.002 | [5 empty] | 52
| 1 | 6 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | 0.002
| Mean |  |  |  | 0.007 |  | 0.002 | 0.005 | [6 empty] | 0.0026667
| Median |  |  |  | 0.007 |  | 0.002 | 0.005 | [6 empty] | 0.0025
| Std Dev. | [5 empty] | 0 | [7 empty] | 0.0008165
| Rel.Std.Dev. | [5 empty] | 0% | [7 empty] | 30.62%
PDM3
| Z-Score (Absolute) |  | 0.67 | 47.88 | 0.51 | 0.67 | 0 | 3.2 | 0.34 | 0.67 | 47.88 | 0.67 | 0.67 | 0.67 | 2.7
| Analytical results for S in OREAS 254c (Certified Value 0.02 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0198 | 0.027967 | 0.46656 | 0.018 | 0.02 | 0.02 | 0.02 | 0.03 | < 0.1 | 0.01812 | 0.02 | 0.020315 | 0.02 | 0.01 | 0.03 | 0.02 | 0.021 |  | 1
| 1 | 2 | 0.0199 | 0.022707 | 0.45502 | 0.018 | 0.02 | 0.02 | 0.02 | 0.03 | < 0.1 | 0.019274 | 0.02 | 0.020652 | 0.02 | 0.01 | 0.03 | 0.02 | 0.022 |  | 18
| 1 | 3 | 0.0199 | 0.027261 | 0.42593 | 0.018 | 0.02 | 0.02 | 0.02 | 0.03 | < 0.1 | 0.017927 | 0.02 | 0.020633 | 0.02 | 0.01 | 0.03 | 0.02 | 0.022 |  | 16
| 1 | 4 | 0.0193 | 0.023455 | 0.43394 | 0.018 | 0.02 | 0.02 | 0.02 | 0.04 | < 0.1 | 0.018537 | 0.02 | 0.019764 | 0.02 | 0.01 | 0.02 | 0.02 | 0.021 |  | 0.019788
| 1 | 5 | 0.0191 | 0.027861 | 0.43576 | 0.017 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.1 | 0.018233 | 0.02 | 0.02028 | 0.02 | 0.01 | 0.03 | 0.02 | 0.021 |  | 53
| 1 | 6 | 0.0192 | 0.024964 | 0.44577 | 0.018 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.1 | 0.01927 | 0.02 | 0.019789 | 0.02 | 0.01 | 0.03 | 0.02 | 0.022
| Mean |  | 0.019533 | 0.025703 | 0.44383 | 0.017833 | 0.02 | 0.02 | 0.02 | 0.028333 |  | 0.01856 | 0.02 | 0.020239 | 0.02 | 0.01 | 0.028333 | 0.02 | 0.0215
| Median |  | 0.01955 | 0.026113 | 0.44076 | 0.018 | 0.02 | 0.02 | 0.02 | 0.03 |  | 0.018385 | 0.02 | 0.020297 | 0.02 | 0.01 | 0.03 | 0.02 | 0.0215
| Std Dev. |  | 0.00037238 | 0.0023145 | 0.015027 | 0.00040825 | 0 | 0 | 0 | 0.0075277 |  | 0.00058576 | 0 | 0.00039022 | 0 | 0 | 0.0040825 | 0 | 0.00054772
| Rel.Std.Dev. |  | 1.906% | 9.005% | 3.386% | 2.289% | 0% | 0% | 0% | 26.57% |  | 3.156% | 0% | 1.928% | 0% | 0% | 14.41% | 0% | 2.548%
| PDM3 |  | -1.286% | 29.89% | 2143% | -9.877% | 1.073% | 1.073% | 1.073% | 43.19% |  | -6.204% | 1.073% | 2.279% | 1.073% | -49.46% | 43.19% | 1.073% | 8.653%
| Z-Score (Absolute) |  | 0.22 | 2.67 | 198.47 | 1.01 | 0 | 0 | 0 | 3.9 | 14.05 | 0.67 | 0 | 0.11 | 0 | 4.68 | 3.9 | 0 | 0.7
| Analytical results for Sb in OREAS 254c (Certified Value 1.17 ppm) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.21 | < 2.5 | 2.1591 | 1.2 | < 0.1 | 1.07 | < 5 | 1.22 | 1.2 | 1.2078 | 1.12 | 1.1414 | 1.1 | 1.08 | 1.24 | 1.08 | 1.21 |  | 1
| 1 | 2 | 1.14 | < 2.5 | < 0.1 | 1.2 | < 0.1 | 1.12 | < 5 | 1.21 | 1.2 | 1.2288 | 1.12 | 1.1436 | 1.16 | 1.11 | 1.2 | 1.12 | 1.2 |  | 34
| 1 | 3 | 1.26 | < 2.5 | 2.7081 | 1.1 | < 0.1 | 1.09 | < 5 | 1.25 | 1.2 | 1.2225 | 1.36 | 1.1723 | 1.13 | 1.1 | 1.2 | 1.02 | 1.3 |  | 16
| 1 | 4 | 1.27 | < 2.5 | 2.5389 | 1.2 | < 0.1 | 1.11 | < 5 | 1.24 | 1.1 | 1.2005 | 1.07 | 0.98068 | 1.1 | 1.09 | 1.2 | 1.15 | 1.22 |  | 1.1652
| 1 | 5 | 1.17 | < 2.5 | 5.4968 | 1.2 | < 0.1 | 1.08 | < 5 | 1.19 | 1.2 | 1.2026 | 1.22 | 1.2323 | 1.13 | 1.08 | 1.22 | 1.12 | 1.25 |  | 54
| 1 | 6 | 1.13 | < 2.5 | 4.0706 | 1.2 | < 0.1 | 1.07 | < 5 | 1.26 | 1.2 | 1.2876 | 1.18 | 0.98592 | 1.14 | 1.12 | 1.2 | 1.08 | 1.28
| Mean |  | 1.1967 |  | 3.3947 | 1.1833 |  | 1.09 |  | 1.2283 | 1.1833 | 1.225 | 1.1783 | 1.1094 | 1.1267 | 1.0967 | 1.21 | 1.095 | 1.2433
| Median |  | 1.19 |  | 2.7081 | 1.2 |  | 1.085 |  | 1.23 | 1.2 | 1.2151 | 1.15 | 1.1425 | 1.13 | 1.095 | 1.2 | 1.1 | 1.235
| Std Dev. |  | 0.059889 |  | 1.379 | 0.040825 |  | 0.020976 |  | 0.026394 | 0.040825 | 0.032672 | 0.10323 | 0.10303 | 0.023381 | 0.01633 | 0.016733 | 0.045497 | 0.040332
| Rel.Std.Dev. |  | 5.005% |  | 40.62% | 3.45% |  | 1.924% |  | 2.149% | 3.45% | 2.667% | 8.761% | 9.287% | 2.075% | 1.489% | 1.383% | 4.155% | 3.244%
| PDM3 |  | 2.701% |  | 191.3% | 1.557% |  | -6.453% |  | 5.419% | 1.557% | 5.13% | 1.128% | -4.79% | -3.306% | -5.881% | 3.846% | -6.024% | 6.706%
| Z-Score (Absolute) |  | 0.15 | 0.75 | 18.59 | 0 | 12.74 | 1.05 | 14.8 | 0.51 | 0 | 0.47 | 0.06 | 0.83 | 0.64 | 0.97 | 0.3 | 0.99 | 0.67
| Analytical results for Sc in OREAS 254c (Certified Value 10.1 ppm) | [19 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 10.2 | 10.286 | 9 | 11 | 10.4 | 9.5756 | 10.4 | 10 | 9.5841 | 8.5 | 9.9854 | 9.8 | 10.4 | 10.5 | 9.7 | 11.2 |  |  | 1
| 1 | 2 | 10.1 | 10.2243 | 9 | 11 | 10.6 | 9.5682 | 10.2 | 10 | 9.5537 | 8.2 | 10.0991 | 9.9 | 10.2 | 10.4 | 9.8 | 11.5 |  |  | 35
| 1 | 3 | 9.8 | 10.1503 | 9 | 11 | 10.7 | 9.8309 | 10.1 | 10 | 9.5347 | 8.1 | 10.0483 | 10 | 10.3 | 10.6 | 8.8 | 11.5 |  |  | 16
| 1 | 4 | 10 | 10.2682 | 9 | 11 | 11 | 9.768 | 10 | 10 | 9.4842 | 8 | 10.0682 | 10.2 | 10.4 | 10.6 | 10.2 | 11.3 |  |  | 10.1107
| 1 | 5 | 9.9 | 10.3066 | 9 | 11 | 10.9 | 9.7754 | 10 | 10 | 9.5146 | 9.1 | 10.1043 | 9.9 | 10.4 | 10.6 | 9.9 | 11.8 |  |  | 55
| 1 | 6 | 10 | 10.3759 | 9 | 11 | 10.5 | 9.9493 | 10.2 | 10 | 9.5839 | 8.8 | 10.1043 | 10.1 | 10.5 | 10.6 | 9.7 | 11.9
| Mean |  | 10 | 10.2685 | 9 | 11 | 10.6833 | 9.7446 | 10.15 | 10 | 9.5425 | 8.45 | 10.0683 | 9.9833 | 10.3667 | 10.55 | 9.6833 | 11.5333
| Median |  | 10 | 10.2771 | 9 | 11 | 10.65 | 9.7717 | 10.15 | 10 | 9.5442 | 8.35 | 10.0837 | 9.95 | 10.4 | 10.6 | 9.75 | 11.5
| Std Dev. |  | 0.14142 | 0.076437 | 0 | 0 | 0.23166 | 0.14868 | 0.15166 | 0 | 0.03953 | 0.43243 | 0.046518 | 0.1472 | 0.10328 | 0.083666 | 0.47081 | 0.27325
| Rel.Std.Dev. |  | 1.414% | 0.7444% | 0% | 0% | 2.168% | 1.526% | 1.494% | 0% | 0.4142% | 5.118% | 0.462% | 1.474% | 0.9963% | 0.793% | 4.862% | 2.369%
| PDM3 |  | -1.094% | 1.562% | -10.99% | 8.796% | 5.664% | -3.621% | 0.3891% | -1.094% | -5.619% | -16.42% | -0.4191% | -1.259% | 2.532% | 4.345% | -4.226% | 14.07%
| Z-Score (Absolute) |  | 0.14 | 0.42 | NA | NA | 1.28 | 0.67 | 0.17 | NA | 1.1 | 3.37 | 0 | 0.18 | 0.62 | 1 | 0.8 | 3.05
Results from laboratories 1.07, 1.09 and 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 254c (Indicative Value 1 ppm) | [19 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.7 | < 1.5 | 13.7653 | < 1 | 0.6 | < 1 | < 5 | 3 | < 1 | 0.072 | < 1 | < 2 | 1.3 | < 1 | < 1 | < 1 |  |  | 1
| 1 | 2 | 1 | < 1.5 | 14.0566 | < 1 | 0.7 | < 1 | < 5 | 2 | < 1 | 0.1453 | < 1 | < 2 | 1.4 | < 1 | < 1 | < 1 |  |  | 1
| 1 | 3 | 0.8 | < 1.5 | 13.6016 | < 1 | 0.7 | < 1 | < 5 | 3 | < 1 | 0.1473 | < 1 | < 2 | 1.3 | < 1 | < 1 | < 1 |  |  | 16
| 1 | 4 | 1 | < 1.5 | 13.672 | < 1 | 0.5 | < 1 | < 5 | 1 | 1 | 0.1571 | < 1 | < 2 | 1.4 | < 1 | < 1 | 1 |  |  | 0.99664
| 1 | 5 | 0.8 | < 1.5 | 12.7395 | < 1 | 0.6 | < 1 | < 5 | 1 | < 1 | 0.1744 | < 1 | < 2 | 1.2 | < 1 | < 1 | < 1 |  |  | 7
| 1 | 6 | 0.9 | < 1.5 | 18.8272 | < 1 | 0.5 | < 1 | < 5 | 2 | < 1 | 0.1751 | < 1 | < 2 | 1.5 | < 1 | < 1 | < 1
| Mean |  | 0.86667 |  | 14.4437 |  | 0.6 |  |  | 2 | 1 | 0.1452 |  |  | 1.35 |  |  | 1
| Median |  | 0.85 |  | 13.7187 |  | 0.6 |  |  | 2 | 1 | 0.1522 |  |  | 1.35 |  |  | 1
| Std Dev. |  | 0.12111 |  | 2.1924 |  | 0.089443 |  |  | 0.89443 |  | 0.038081 |  |  | 0.10488
| Rel.Std.Dev. |  | 13.97% |  | 15.18% |  | 14.91% |  |  | 44.72% |  | 26.23% |  |  | 7.769%
| PDM3 |  | -13.04% |  | 1349% |  | -39.8% |  |  | 100.7% | 0.3367% | -85.43% |  |  | 35.45% |  |  | 0.3367%
| Z-Score (Absolute) |  | 1.48 | 0.85 | 74.72 | 0.49 | 0.04 | 0.49 | 10.29 | 7.6 | 0.04 | 2.41 | 0.49 | 2.2 | 4.09 | 0.49 | 0.49 | 0.04
| Analytical results for Sm in OREAS 254c (Certified Value 6.04 ppm) | [19 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 6.27 | 6.0299 | 6.962 | 6.1 | 5.92 | 5.9603 | 6.74 | [11 empty] | 1
| 1 | 2 | 5.87 | 6.2156 | 6.887 | 7 | 6.03 | 6.0153 | 7.11 | [11 empty] | 5
| 1 | 3 | 5.97 | 6.2279 | 6.4709 | 5.2 | 5.98 | 5.8754 | 7.04 | [11 empty] | 16
| 1 | 4 | 6.18 | 6.0135 | 6.6282 | 4.9 | 6.12 | 6.0411 | 6.97 | [11 empty] | 6.0395
| 1 | 5 | 5.93 | 6.2948 | 6.7411 | 5.4 | 6.14 | 5.9008 | 7.13 | [11 empty] | 56
| 1 | 6 | 6.12 | 6.3508 | 6.656 | 6.1 | 5.92 | 6.0491 | 7.09
| Mean |  | 6.0567 | 6.1888 | 6.7242 | 5.7833 | 6.0183 | 5.9737 | 7.0133
| Median |  | 6.045 | 6.2218 | 6.6986 | 5.75 | 6.005 | 5.9878 | 7.065
| Std Dev. |  | 0.15693 | 0.13837 | 0.17965 | 0.7679 | 0.096003 | 0.073633 | 0.14569
| Rel.Std.Dev. |  | 2.591% | 2.236% | 2.672% | 13.28% | 1.595% | 1.233% | 2.077%
| PDM3 |  | 0.2844% | 2.472% | 11.34% | -4.241% | -0.3503% | -1.09% | 16.12%
| Z-Score (Absolute) |  | 0 | 0.67 | 3.41 | 1.4 | 0.2 | 0.42 | 4.88
| Analytical results for Sn in OREAS 254c (Certified Value 3.75 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 3.9 | 3.821 | 38.9202 | 3.7 | < 1 | 3.5 | < 10 | 4.2 | 3.9 | 3.6914 | 3.5 | 3.9039 | 3.7 | 3.6 | 3.8 | 3.5 | 3.8 |  | 1
| 1 | 2 | 3.8 | 3.6805 | 37.9323 | 4.1 | 1 | 3.6 | < 10 | 3.9 | 3.7 | 3.5884 | 3.4 | 3.8431 | 3.7 | 3.5 | 3.7 | 3.7 | 3.9 |  | 21
| 1 | 3 | 3.9 | 4.1922 | 40.625 | 3.8 | < 1 | 3.5 | < 10 | 3.8 | 3.8 | 3.5944 | 3.6 | 3.7568 | 3.8 | 3.5 | 3.8 | 3.1 | 4 |  | 16
| 1 | 4 | 3.8 | 3.662 | 39.7486 | 4 | < 1 | 3.7 | < 10 | 3.8 | 3.8 | 3.4674 | 3.3 | 3.8274 | 3.9 | 3.5 | 3.8 | 3.8 | 3.8 |  | 3.748
| 1 | 5 | 3.8 | 3.9049 | 39.9068 | 3.9 | < 1 | 3.6 | < 10 | 3.6 | 4 | 3.7018 | 4 | 3.618 | 3.8 | 3.5 | 3.7 | 3.8 | 3.9 |  | 57
| 1 | 6 | 3.8 | 4.0577 | 41.5681 | 3.8 | < 1 | 3.5 | < 10 | 3.8 | 4.1 | 3.7209 | 3.7 | 3.618 | 3.8 | 3.6 | 3.7 | 3.7 | 4
| Mean |  | 3.8333 | 3.8864 | 39.7835 | 3.8833 | 1 | 3.5667 |  | 3.85 | 3.8833 | 3.6274 | 3.5833 | 3.7612 | 3.7833 | 3.5333 | 3.75 | 3.6 | 3.9
| Median |  | 3.8 | 3.8629 | 39.8277 | 3.85 | 1 | 3.55 |  | 3.8 | 3.85 | 3.6429 | 3.55 | 3.7921 | 3.8 | 3.5 | 3.75 | 3.7 | 3.9
| Std Dev. |  | 0.05164 | 0.20988 | 1.2713 | 0.1472 |  | 0.08165 |  | 0.19748 | 0.1472 | 0.096523 | 0.24833 | 0.12043 | 0.075277 | 0.05164 | 0.054772 | 0.26833 | 0.089443
| Rel.Std.Dev. |  | 1.347% | 5.4% | 3.196% | 3.79% |  | 2.289% |  | 5.129% | 3.79% | 2.661% | 6.93% | 3.202% | 1.99% | 1.462% | 1.461% | 7.454% | 2.293%
| PDM3 |  | 2.278% | 3.693% | 961.5% | 3.612% | -73.32% | -4.837% |  | 2.722% | 3.612% | -3.217% | -4.393% | 0.3526% | 0.9435% | -5.727% | 0.05411% | -3.948% | 4.056%
| Z-Score (Absolute) |  | 0.29 | 0.6 | 208.07 | 0.58 | 18.5 | 1.25 | 7.03 | 0.39 | 0.58 | 0.9 | 1.16 | 0.13 | 0 | 1.44 | 0.19 | 1.06 | 0.67
| Analytical results for Sr in OREAS 254c (Certified Value 285 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 277.98 | 312.9948 | 287.7321 | 412.6899 | 272 | 270 | 286 | 300 | 297 | 299 | 275.701 | 281 | 288.5944 | 288 | 287 | 299 | 279 | 268 | 1
| 1 | 2 | 274.83 | 313.1028 | 284.2952 | 410.4764 | 274 | 280 | 282.7 | 296 | 289 | 299 | 269.5839 | 288 | 290.3722 | 283 | 291 | 297 | 283 | 268 | 22
| 1 | 3 | 270.71 | 311.2452 | 286.231 | 403.4552 | 267 | 278 | 284.5 | 306 | 285 | 304 | 262.1811 | 284 | 288.3771 | 290 | 282 | 299 | 277 | 276 | 16
| 1 | 4 | 274.53 | 303.318 | 284.6095 | 399.4414 | 270 | 271 | 283.8 | 303 | 281 | 296 | 256.2081 | 289 | 272.6173 | 298 | 276 | 239 | 287 | 261 | 285.2934
| 1 | 5 | 268.28 | 312.5088 | 290.3841 | 400.5089 | 272 | 278 | 283.3 | 302 | 274 | 303 | 275.7082 | 294 | 285.6309 | 290 | 284 | 300 | 283 | 264 | 58
| 1 | 6 | 271.71 | 310.77 | 287.531 | 411.2852 | 271 | 279 | 283.5 | 305 | 278 | 291 | 277.2554 | 285 | 271.5351 | 293 | 292 | 298 | 281 | 276
| Mean |  | 273.0067 | 310.6566 | 286.7971 | 406.3095 | 271 | 276 | 283.9667 | 302 | 284 | 298.6667 | 269.4396 | 286.8333 | 282.8545 | 290.3333 | 285.3333 | 288.6667 | 281.6667 | 268.8333
| Median |  | 273.12 | 311.877 | 286.881 | 406.9658 | 271.5 | 278 | 283.65 | 302.5 | 283 | 299 | 272.6425 | 286.5 | 287.004 | 290 | 285.5 | 298.5 | 282 | 268
| Std Dev. |  | 3.4548 | 3.7174 | 2.2647 | 5.8617 | 2.3664 | 4.3359 | 1.1587 | 3.6332 | 8.2462 | 4.761 | 8.5717 | 4.535 | 8.4923 | 5.0067 | 5.9889 | 24.353 | 3.5024 | 6.1455
| Rel.Std.Dev. |  | 1.265% | 1.197% | 0.7897% | 1.443% | 0.8732% | 1.571% | 0.4081% | 1.203% | 2.904% | 1.594% | 3.181% | 1.581% | 3.002% | 1.724% | 2.099% | 8.436% | 1.243% | 2.286%
| PDM3 |  | -4.307% | 8.89% | 0.5271% | 42.42% | -5.01% | -3.257% | -0.465% | 5.856% | -0.4534% | 4.688% | -5.557% | 0.5398% | -0.8549% | 1.767% | 0.014% | 1.182% | -1.271% | -5.77%
| Z-Score (Absolute) |  | 1.1 | 2.45 | 0.2 | 11.45 | 1.29 | 0.82 | 0.07 | 1.63 | 0.06 | 1.32 | 1.43 | 0.2 | 0.17 | 0.53 | 0.06 | 0.38 | 0.28 | 1.49
| Analytical results for Ta in OREAS 254c (Certified Value 1.97 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.19 | 2.5054 | 1.81 | < 0.1 | < 0.05 | 7.1833 | 2.14 | 2 | 2.7946 | 1.79 | 2.0158 | 1.8 | 2 | 2.26 | 2.03 | 1.78 |  |  | 1
| 1 | 2 | 2.1 | 2.5049 | 1.8 | < 0.1 | < 0.05 | 6.52 | 2.11 | 1.9 | 2.8417 | 1.75 | 1.9719 | 1.8 | 2.03 | 2.19 | 1.96 | 1.96 |  |  | 23
| 1 | 3 | 2.12 | 2.4038 | 1.83 | < 0.1 | < 0.05 | 7.0533 | 2.04 | 2 | 2.777 | 1.68 | 1.9867 | 1.7 | 1.99 | 2.24 | 1.73 | 1.88 |  |  | 16
| 1 | 4 | 2.12 | 2.7556 | 1.88 | < 0.1 | < 0.05 | 7.8867 | 2.02 | 1.9 | 2.8034 | 1.69 | 1.8632 | 1.7 | 1.93 | 2.25 | 2.05 | 1.94 |  |  | 1.9663
| 1 | 5 | 2.12 | 2.6518 | 1.81 | < 0.1 | < 0.05 | 7.12 | 2.08 | 2 | 2.7881 | 2.02 | 1.9456 | 1.8 | 1.97 | 2.23 | 2 | 1.85 |  |  | 59
| 1 | 6 | 2.12 | 2.6321 | 1.83 | < 0.1 | < 0.05 | 7.0033 | 2.16 | 2 | 2.8192 | 1.91 | 1.8733 | 1.8 | 2.04 | 2.18 | 1.98 | 1.93
| Mean |  | 2.1283 | 2.5756 | 1.8267 |  |  | 7.1278 | 2.0917 | 1.9667 | 2.804 | 1.8067 | 1.9428 | 1.7667 | 1.9933 | 2.225 | 1.9583 | 1.89
| Median |  | 2.12 | 2.5687 | 1.82 |  |  | 7.0867 | 2.095 | 2 | 2.799 | 1.77 | 1.9588 | 1.8 | 1.995 | 2.235 | 1.99 | 1.905
| Std Dev. |  | 0.031252 | 0.12716 | 0.028752 |  |  | 0.44034 | 0.055287 | 0.05164 | 0.023336 | 0.13367 | 0.062103 | 0.05164 | 0.040332 | 0.032711 | 0.11652 | 0.067528
| Rel.Std.Dev. |  | 1.468% | 4.937% | 1.574% |  |  | 6.178% | 2.643% | 2.626% | 0.8322% | 7.399% | 3.197% | 2.923% | 2.023% | 1.47% | 5.95% | 3.573%
| PDM3 |  | 8.24% | 30.99% | -7.102% |  |  | 262.5% | 6.375% | 0.01797% | 42.6% | -8.119% | -1.197% | -10.15% | 1.374% | 13.16% | -0.4058% | -3.881%
| Z-Score (Absolute) |  | 0.7 | 2.57 | 0.57 | 8.02 | 8.12 | 21.66 | 0.54 | 0.02 | 3.53 | 0.65 | 0.08 | 0.82 | 0.13 | 1.1 | 0.02 | 0.3
| Analytical results for Tb in OREAS 254c (Certified Value 0.68 ppm) | [19 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.17 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.68 | 0.72952 | 0.6877 | 0.6 | 0.67 | 0.53935 | 0.7 | 0.65 | [10 empty] | 1
| 1 | 2 | 0.66 | 0.74459 | 0.71314 | 0.6 | 0.67 | 0.5655 | 0.7 | 0.65 | [10 empty] | 24
| 1 | 3 | 0.67 | 0.7711 | 0.66245 | 0.6 | 0.68 | 0.49181 | 0.7 | 0.66 | [10 empty] | 16
| 1 | 4 | 0.67 | 0.7211 | 0.64052 | 0.7 | 0.68 | 0.54506 | 0.7 | 0.62 | [10 empty] | 0.6762
| 1 | 5 | 0.69 | 0.73356 | 0.67373 | 0.6 | 0.67 | 0.53598 | 0.7 | 0.66 | [10 empty] | 60
| 1 | 6 | 0.65 | 0.74618 | 0.68063 | 0.6 | 0.67 | 0.58207 | 0.7 | 0.66
| Mean |  | 0.67 | 0.74101 | 0.67636 | 0.61667 | 0.67333 | 0.54329 | 0.7 | 0.65
| Median |  | 0.67 | 0.73908 | 0.67718 | 0.6 | 0.67 | 0.5422 | 0.7 | 0.655
| Std Dev. |  | 0.014142 | 0.017487 | 0.02443 | 0.040825 | 0.005164 | 0.030723 | 1.2162e-16 | 0.015492
| Rel.Std.Dev. |  | 2.111% | 2.36% | 3.612% | 6.62% | 0.7669% | 5.655% | 1.737e-14% | 2.383%
| PDM3 |  | -0.9163% | 9.585% | 0.02448% | -8.804% | -0.4233% | -19.65% | 3.52% | -3.874%
| Z-Score (Absolute) |  | 0.04 | 1.87 | 0.13 | 1.48 | 0.04 | 3.46 | 0.76 | 0.58
| Analytical results for Te in OREAS 254c (Indicative Value 0.071 ppm) | [19 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | < 0.2 | 0.10634 | < 0.1 | < 0.1 | 0.08 | 6.888 | < 0.05 | < 0.5 | 0.053 | < 0.05 | 0.15708 | 0.03 | < 0.05 | 0.06 | < 0.05 |  |  |  | 1
| 1 | 2 | < 0.2 | 0.11105 | < 0.1 | < 0.1 | 0.08 | 7.1707 | < 0.05 | < 0.5 | 0.0482 | < 0.05 | 0.17867 | 0.12 | 0.07 | 0.07 | < 0.05 |  |  |  | 2
| 1 | 3 | < 0.2 | < 0.1 | < 0.1 | < 0.1 | 0.07 | 7.5867 | < 0.05 | < 0.5 | 0.0626 | 0.05 | 0.19826 | 0.08 | 0.07 | 0.07 | < 0.05 |  |  |  | 16
| 1 | 4 | < 0.2 | 0.11242 | < 0.1 | < 0.1 | 0.08 | 7.7787 | 0.05 | < 0.5 | 0.0434 | 0.05 | 0.20301 | 0.05 | 0.05 | 0.06 | < 0.05 |  |  |  | 0.071049
| 1 | 5 | < 0.2 | 0.11096 | < 0.1 | < 0.1 | 0.08 | 7.808 | < 0.05 | < 0.5 | 0.0578 | 0.08 | 0.18279 | 0.12 | 0.07 | 0.07 | < 0.05 |  |  |  | 8
| 1 | 6 | < 0.2 | 0.1398 | < 0.1 | < 0.1 | 0.08 | 7.568 | < 0.05 | < 0.5 | 0.0482 | 0.08 | 0.15797 | 0.09 | 0.07 | 0.06 | < 0.05
| Mean |  |  | 0.11611 |  |  | 0.078333 | 7.4667 | 0.05 |  | 0.0522 | 0.065 | 0.17963 | 0.081667 | 0.066 | 0.065
| Median |  |  | 0.11105 |  |  | 0.08 | 7.5773 | 0.05 |  | 0.0506 | 0.065 | 0.18073 | 0.085 | 0.07 | 0.065
| Std Dev. |  |  | 0.013438 |  |  | 0.0040825 | 0.36359 |  |  | 0.0070654 | 0.017321 | 0.019403 | 0.03656 | 0.0089443 | 0.0054772
| Rel.Std.Dev. |  |  | 11.57% |  |  | 5.212% | 4.87% |  |  | 13.54% | 26.65% | 10.8% | 44.77% | 13.55% | 8.427%
| PDM3 |  |  | 63.43% |  |  | 10.25% | 10409% | -29.63% |  | -26.53% | -8.514% | 152.8% | 14.94% | -7.107% | -8.514%
| Z-Score (Absolute) |  | 1.42 | 1.62 | 0.61 | 0.61 | 0.54 | 299.46 | 1.45 | 7.48 | 0.52 | 0.54 | 4.64 | 0.67 | 0.24 | 0 | 1.62
| Analytical results for Th in OREAS 254c (Certified Value 8.37 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 8.66 | 8.2104 | 8.8788 | 8.25 | 8.8 | 7.44 | 8.91 | 9.3 | 11.1399 | 7.69 | 8.6479 | 7.28 | 8.18 | 9.08 | 8.24 | 8.18 |  |  | 1
| 1 | 2 | 8.71 | 8.2098 | 8.8606 | 8.41 | 9.7 | 7.51 | 8.58 | 8.7 | 11.3185 | 7.56 | 8.5012 | 7.38 | 8.08 | 8.97 | 8.22 | 8.28 |  |  | 26
| 1 | 3 | 8.58 | 8.2702 | 9.1581 | 8.29 | 9 | 7.64 | 8.82 | 9.3 | 11.0149 | 7.34 | 8.7333 | 7.55 | 8.04 | 8.83 | 7.18 | 8.38 |  |  | 16
| 1 | 4 | 8.72 | 8.1612 | 8.3731 | 8.2 | 8.8 | 7.65 | 8.38 | 8.5 | 10.757 | 7.34 | 8.0736 | 7 | 7.94 | 8.76 | 8.53 | 8.28 |  |  | 8.3684
| 1 | 5 | 8.77 | 8.389 | 9.1619 | 8.27 | 9 | 7.63 | 8.88 | 8.7 | 11.0306 | 8.56 | 8.3936 | 7.33 | 8.01 | 8.72 | 8.41 | 8.19 |  |  | 61
| 1 | 6 | 8.77 | 8.1154 | 8.6634 | 8.5 | 9.3 | 7.45 | 9.09 | 8.8 | 11.4268 | 8.03 | 8.2088 | 7.47 | 8.31 | 9.03 | 8.26 | 8.35
| Mean |  | 8.7017 | 8.226 | 8.8493 | 8.32 | 9.1 | 7.5533 | 8.7767 | 8.8833 | 11.1146 | 7.7533 | 8.4264 | 7.335 | 8.0933 | 8.8983 | 8.14 | 8.2767
| Median |  | 8.715 | 8.2101 | 8.8697 | 8.28 | 9 | 7.57 | 8.85 | 8.75 | 11.0853 | 7.625 | 8.4474 | 7.355 | 8.06 | 8.9 | 8.25 | 8.28
| Std Dev. |  | 0.072503 | 0.095334 | 0.30181 | 0.11243 | 0.34641 | 0.098116 | 0.25477 | 0.33714 | 0.23856 | 0.47124 | 0.25368 | 0.19066 | 0.13261 | 0.14905 | 0.48526 | 0.081158
| Rel.Std.Dev. |  | 0.8332% | 1.159% | 3.411% | 1.351% | 3.807% | 1.299% | 2.903% | 3.795% | 2.146% | 6.078% | 3.011% | 2.599% | 1.639% | 1.675% | 5.961% | 0.9806%
| PDM3 |  | 3.983% | -1.701% | 5.747% | -0.5779% | 8.743% | -9.739% | 4.879% | 6.154% | 32.82% | -7.349% | 0.6937% | -12.35% | -3.286% | 6.333% | -2.729% | -1.096%
| Z-Score (Absolute) |  | 0.5 | 0.23 | 0.73 | 0.08 | 1.11 | 1.26 | 0.62 | 0.78 | 4.2 | 0.95 | 0.08 | 1.59 | 0.43 | 0.81 | 0.36 | 0.15
| Analytical results for Ti in OREAS 254c (Certified Value 0.532 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.5395 | 0.5725 | 0.52188 | < 0.1 | 0.52 | 0.334 | 0.538 | 0.5 | 0.551 | 0.536 | 0.39832 | 0.514 | 0.55484 | 0.52 | 0.537 | 0.539 | 0.517 | 0.471 | 1
| 1 | 2 | 0.5363 | 0.555 | 0.51645 | 0.4612 | 0.533 | 0.376 | 0.537 | 0.5 | 0.53 | 0.538 | 0.40535 | 0.527 | 0.55413 | 0.52 | 0.535 | 0.54 | 0.514 | 0.469 | 27
| 1 | 3 | 0.5314 | 0.55348 | 0.52107 | < 0.1 | 0.531 | 0.228 | 0.533 | 0.5 | 0.529 | 0.539 | 0.39397 | 0.524 | 0.55707 | 0.52 | 0.521 | 0.536 | 0.514 | 0.485 | 16
| 1 | 4 | 0.5291 | 0.56595 | 0.52566 | 0.45944 | 0.523 | 0.229 | 0.538 | 0.51 | 0.521 | 0.541 | 0.40436 | 0.526 | 0.55318 | 0.54 | 0.516 | 0.382 | 0.519 | 0.46 | 0.53161
| 1 | 5 | 0.5333 | 0.56298 | 0.52744 | 0.46727 | 0.531 | 0.201 | 0.54 | 0.51 | 0.529 | 0.548 | 0.39356 | 0.534 | 0.55656 | 0.53 | 0.535 | 0.533 | 0.516 | 0.471 | 62
| 1 | 6 | 0.5341 | 0.55291 | 0.53 | 0.47812 | 0.532 | 0.274 | 0.533 | 0.51 | 0.536 | 0.53 | 0.40503 | 0.519 | 0.55148 | 0.52 | 0.547 | 0.532 | 0.513 | 0.488
| Mean |  | 0.53395 | 0.56047 | 0.52375 | 0.46651 | 0.52833 | 0.27367 | 0.5365 | 0.505 | 0.53267 | 0.53867 | 0.4001 | 0.524 | 0.55454 | 0.525 | 0.53183 | 0.51033 | 0.5155 | 0.474
| Median |  | 0.5337 | 0.55899 | 0.52377 | 0.46424 | 0.531 | 0.2515 | 0.5375 | 0.505 | 0.5295 | 0.5385 | 0.40134 | 0.525 | 0.55448 | 0.52 | 0.535 | 0.5345 | 0.515 | 0.471
| Std Dev. |  | 0.003655 | 0.0079634 | 0.0049047 | 0.0084366 | 0.0054283 | 0.068518 | 0.002881 | 0.0054772 | 0.010172 | 0.0059217 | 0.0055412 | 0.0068993 | 0.0020944 | 0.0083666 | 0.011356 | 0.06295 | 0.0022583 | 0.010545
| Rel.Std.Dev. |  | 0.6845% | 1.421% | 0.9365% | 1.808% | 1.027% | 25.04% | 0.537% | 1.085% | 1.91% | 1.099% | 1.385% | 1.317% | 0.3777% | 1.594% | 2.135% | 12.34% | 0.4381% | 2.225%
| PDM3 |  | 0.4401% | 5.429% | -1.479% | -12.25% | -0.6165% | -48.52% | 0.9197% | -5.006% | 0.1987% | 1.327% | -24.74% | -1.432% | 4.314% | -1.243% | 0.04191% | -4.002% | -3.031% | -10.84%
| Z-Score (Absolute) |  | 0.49 | 1.85 | 0.04 | 10.14 | 0.2 | 12.93 | 0.62 | 1.01 | 0.42 | 0.73 | 6.41 | 0.03 | 1.55 | 0.03 | 0.38 | 0.73 | 0.46 | 2.6
| Analytical results for Tl in OREAS 254c (Certified Value 0.36 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.38 |  | 0.37798 | 0.37 | 0.36 | < 0.02 | 0.36 | < 0.5 | < 0.6762 | 0.35 | 0.34694 | 0.37 | 0.34 | 0.36 | 0.33 | 0.36 |  |  | 1
| 1 | 2 | 0.38 |  | 0.3598 | 0.38 | 0.38 | < 0.02 | 0.34 | < 0.5 | < 0.6687 | 0.34 | 0.37224 | 0.38 | 0.34 | 0.35 | 0.34 | 0.37 |  |  | 28
| 1 | 3 | 0.38 |  | 0.37868 | 0.37 | 0.36 | < 0.02 | 0.36 | < 0.5 | < 0.6878 | 0.33 | 0.38425 | 0.37 | 0.35 | 0.35 | 0.35 | 0.37 |  |  | 16
| 1 | 4 | 0.38 |  | 0.34894 | 0.37 | 0.37 | < 0.02 | 0.33 | < 0.5 | < 0.7115 | 0.34 | 0.28786 | 0.38 | 0.33 | 0.35 | 0.34 | 0.36 |  |  | 0.36087
| 1 | 5 | 0.39 |  | 0.38166 | 0.38 | 0.37 | < 0.02 | 0.34 | < 0.5 | < 0.6803 | 0.35 | 0.35985 | 0.38 | 0.34 | 0.35 | 0.33 | 0.36 |  |  | 63
| 1 | 6 | 0.38 |  | 0.37083 | 0.38 | 0.39 | < 0.02 | 0.36 | < 0.5 | < 0.6774 | 0.36 | 0.28564 | 0.38 | 0.35 | 0.35 | 0.34 | 0.37
| Mean |  | 0.38167 |  | 0.36965 | 0.375 | 0.37167 |  | 0.34833 |  |  | 0.345 | 0.33946 | 0.37667 | 0.34167 | 0.35167 | 0.33833 | 0.365
| Median |  | 0.38 |  | 0.37441 | 0.375 | 0.37 |  | 0.35 |  |  | 0.345 | 0.35339 | 0.38 | 0.34 | 0.35 | 0.34 | 0.365
| Std Dev. |  | 0.0040825 |  | 0.012826 | 0.0054772 | 0.01169 |  | 0.013292 |  |  | 0.010488 | 0.042686 | 0.005164 | 0.0075277 | 0.0040825 | 0.0075277 | 0.0054772
| Rel.Std.Dev. |  | 1.07% |  | 3.47% | 1.461% | 3.145% |  | 3.816% |  |  | 3.04% | 12.57% | 1.371% | 2.203% | 1.161% | 2.225% | 1.501%
| PDM3 |  | 5.762% |  | 2.432% | 3.915% | 2.991% |  | -3.475% |  |  | -4.398% | -5.933% | 4.377% | -5.322% | -2.551% | -6.246% | 1.144%
| Z-Score (Absolute) |  | 1.35 | NA | 0.86 | 1.08 | 0.94 | 13.69 | 0 | 3.98 | 0.26 | 0.13 | 0.36 | 1.15 | 0.27 | 0.13 | 0.4 | 0.67
| Analytical results for Tm in OREAS 254c (Certified Value 0.2 ppm) | [19 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.2 | 0.21129 | 0.19906 | 0.2 | 0.19 | 0.2013 | 0.17 | [11 empty] | 1
| 1 | 2 | 0.19 | 0.22596 | 0.19542 | 0.2 | 0.19 | 0.2031 | 0.18 | [11 empty] | 7
| 1 | 3 | 0.19 | 0.21129 | 0.19631 | 0.2 | 0.19 | 0.1989 | 0.19 | [11 empty] | 16
| 1 | 4 | 0.19 | 0.20629 | 0.18962 | 0.2 | 0.2 | 0.1957 | 0.18 | [11 empty] | 0.19532
| 1 | 5 | 0.19 | 0.20043 | 0.18452 | 0.2 | 0.2 | 0.2065 | 0.19 | [11 empty] | 64
| 1 | 6 | 0.2 | 0.20336 | 0.18752 | 0.2 | 0.19 | 0.2048 | 0.18
| Mean |  | 0.19333 | 0.20977 | 0.19208 | 0.2 | 0.19333 | 0.20172 | 0.18167
| Median |  | 0.19 | 0.20879 | 0.19252 | 0.2 | 0.19 | 0.2022 | 0.18
| Std Dev. |  | 0.005164 | 0.0090236 | 0.0056881 | 3.0405e-17 | 0.005164 | 0.0039651 | 0.0075277
| Rel.Std.Dev. |  | 2.671% | 4.302% | 2.961% | 1.52e-14% | 2.671% | 1.966% | 4.144%
| PDM3 |  | -1.015% | 7.402% | -1.659% | 2.398% | -1.015% | 3.277% | -6.989%
| Z-Score (Absolute) |  | 0 | 2.3 | 0.18 | NA | 0 | 1.17 | 1.63
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 254c (Certified Value 1.53 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.64 | 1.4997 | 1.5437 | 1.52 | 1.6 | 1.5 | 1.6 | 1.6 | 1.6852 | 1.5 | 1.6597 | 1.55 | 1.4 | 1.6 | 1.5 | 1.35 |  |  | 1
| 1 | 2 | 1.6 | 1.5045 | 1.4974 | 1.53 | 1.8 | 1.5 | 1.6 | 1.5 | 1.7329 | 1.6 | 1.6784 | 1.54 | 1.4 | 1.5 | 1.5 | 1.49 |  |  | 30
| 1 | 3 | 1.62 | 1.5281 | 1.5445 | 1.54 | 1.7 | 1.5 | 1.5 | 1.6 | 1.69 | 1.4 | 1.6043 | 1.56 | 1.4 | 1.6 | 1.5 | 1.45 |  |  | 16
| 1 | 4 | 1.58 | 1.4651 | 1.394 | 1.55 | 1.5 | 1.5 | 1.5 | 1.4 | 1.6693 | 1.4 | 1.4994 | 1.59 | 1.4 | 1.5 | 1.6 | 1.34 |  |  | 1.5321
| 1 | 5 | 1.66 | 1.5194 | 1.542 | 1.56 | 1.7 | 1.6 | 1.6 | 1.5 | 1.6593 | 1.5 | 1.5344 | 1.56 | 1.4 | 1.6 | 1.5 | 1.44 |  |  | 65
| 1 | 6 | 1.54 | 1.4966 | 1.5482 | 1.56 | 1.7 | 1.5 | 1.7 | 1.5 | 1.7507 | 1.5 | 1.4905 | 1.56 | 1.5 | 1.6 | 1.5 | 1.43
| Mean |  | 1.6067 | 1.5022 | 1.5116 | 1.5433 | 1.6667 | 1.5167 | 1.5833 | 1.5167 | 1.6979 | 1.4833 | 1.5778 | 1.56 | 1.4167 | 1.5667 | 1.5167 | 1.4167
| Median |  | 1.61 | 1.5021 | 1.5428 | 1.545 | 1.7 | 1.5 | 1.6 | 1.5 | 1.6876 | 1.5 | 1.5693 | 1.56 | 1.4 | 1.6 | 1.5 | 1.435
| Std Dev. |  | 0.043205 | 0.021849 | 0.060681 | 0.01633 | 0.10328 | 0.040825 | 0.075277 | 0.075277 | 0.036179 | 0.075277 | 0.081469 | 0.016733 | 0.040825 | 0.05164 | 0.040825 | 0.059217
| Rel.Std.Dev. |  | 2.689% | 1.454% | 4.014% | 1.058% | 6.197% | 2.692% | 4.754% | 4.963% | 2.131% | 5.075% | 5.164% | 1.073% | 2.882% | 3.296% | 2.692% | 4.18%
| PDM3 |  | 4.866% | -1.952% | -1.338% | 0.7318% | 8.782% | -1.009% | 3.343% | -1.009% | 10.82% | -3.184% | 2.98% | 1.82% | -7.536% | 2.255% | -1.009% | -7.536%
| Z-Score (Absolute) |  | 1.24 | 0.45 | 0.3 | 0.22 | 2.21 | 0.22 | 0.86 | 0.22 | 2.72 | 0.76 | 0.77 | 0.49 | 1.83 | 0.59 | 0.22 | 1.83
| Analytical results for V in OREAS 254c (Certified Value 69 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 69 | 67.1202 | 70.3913 | 57.5212 | 68 | 42 | 68 | 70.7233 | 71 | 76 | 61.4605 | 68 | 71.9769 | 68 | 71 | 71 | 69 | 67 | 1
| 1 | 2 | 70 | 65.758 | 68.7725 | 57.0568 | 67 | 51 | 68 | 70.3733 | 69 | 76 | 61.0866 | 70 | 72.5471 | 66 | 71 | 71 | 69 | 67 | 31
| 1 | 3 | 71 | 65.6306 | 70.2094 | 57.7541 | 68 | 34 | 68 | 72.0667 | 70 | 76 | 60.8569 | 69 | 73.3141 | 68 | 69 | 70 | 68 | 67 | 16
| 1 | 4 | 68 | 65.1602 | 68.9519 | 57.1415 | 67 | 35 | 70 | 72.3167 | 68 | 75 | 60.9683 | 70 | 72.7428 | 70 | 69 | 57 | 70 | 64 | 69.3031
| 1 | 5 | 70 | 64.974 | 70.4075 | 56.878 | 68 | 31 | 69 | 72.18 | 70 | 77 | 60.3203 | 71 | 70.7625 | 68 | 70 | 71 | 70 | 67 | 66
| 1 | 6 | 70 | 67.1496 | 70.3337 | 58.5694 | 38 | 38 | 68 | 72.3633 | 69 | 73 | 61.0617 | 70 | 71.6378 | 69 | 72 | 70 | 69 | 69
| Mean |  | 69.6667 | 65.9654 | 69.8444 | 57.4868 | 62.6667 | 38.5 | 68.5 | 71.6706 | 69.5 | 75.5 | 60.959 | 69.6667 | 72.1635 | 68.1667 | 70.3333 | 68.3333 | 69.1667 | 66.8333
| Median |  | 70 | 65.6943 | 70.2716 | 57.3314 | 67.5 | 36.5 | 68 | 72.1233 | 69.5 | 76 | 61.015 | 70 | 72.262 | 68 | 70.5 | 70.5 | 69 | 67
| Std Dev. |  | 1.0328 | 0.95102 | 0.76605 | 0.61958 | 12.0941 | 7.1764 | 0.83666 | 0.88248 | 1.0488 | 1.3784 | 0.3733 | 1.0328 | 0.9034 | 1.3292 | 1.2111 | 5.5737 | 0.75277 | 1.6021
| Rel.Std.Dev. |  | 1.482% | 1.442% | 1.097% | 1.078% | 19.3% | 18.64% | 1.221% | 1.231% | 1.509% | 1.826% | 0.6124% | 1.482% | 1.252% | 1.95% | 1.722% | 8.157% | 1.088% | 2.397%
| PDM3 |  | 0.5245% | -4.816% | 0.781% | -17.05% | -9.576% | -44.45% | -1.159% | 3.416% | 0.2841% | 8.942% | -12.04% | 0.5245% | 4.127% | -1.64% | 1.487% | -1.399% | -0.1969% | -3.564%
| Z-Score (Absolute) |  | 0.32 | 1.11 | 0.39 | 4.37 | 2.38 | 11.69 | 0.13 | 1.09 | 0.26 | 2.57 | 3.03 | 0.32 | 1.28 | 0.26 | 0.58 | 0.19 | 0.13 | 0.77
| Analytical results for W in OREAS 254c (Certified Value 3.2 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 3.6 | 3.753 | 3.2509 | 5.9932 | 3 | 0.1 | 2.1 | < 10 | 3.2 | 3.1 | 3.4714 | 3.1 | 3.1223 | 3.3 | 3 | 3.4 | 3.4 | 3.2 | 1
| 1 | 2 | 3.3 | 3.7182 | 3.0717 | 7.3315 | 3 | 0.2 | 2.2 | < 10 | 3.2 | 3 | 3.6096 | 2.9 | 3.0407 | 3.2 | 3.2 | 3.4 | 3.3 | 3.2 | 32
| 1 | 3 | 3.4 | 3.8934 | 3.2864 | 6.1666 | 3 | < 0.1 | 2.2 | < 10 | 3.1 | 3 | 3.5062 | 3 | 3.1554 | 3.2 | 3 | 3.5 | 2.7 | 3.1 | 16
| 1 | 4 | 3.3 | 3.8576 | 3.0831 | 6.608 | 3.1 | 0.2 | 2.2 | < 10 | 3.2 | 3 | 3.457 | 2.9 | 3.1479 | 3.3 | 3 | 3.4 | 3.2 | 3.3 | 3.203
| 1 | 5 | 3.3 | 3.8731 | 3.3706 | 6.779 | 3 | 0.1 | 2.2 | < 10 | 3 | 3 | 3.5661 | 3.3 | 3.0217 | 3.2 | 3 | 3.4 | 3.3 | 3.1 | 67
| 1 | 6 | 3.3 | 3.3659 | 3.1409 | 7.2986 | 3.2 | < 0.1 | 2.1 | < 10 | 3.3 | 3.1 | 3.7309 | 3.2 | 2.9433 | 3.2 | 3.3 | 3.4 | 3.3 | 3.2
| Mean |  | 3.3667 | 3.7435 | 3.2006 | 6.6961 | 3.05 | 0.15 | 2.1667 |  | 3.1667 | 3.0333 | 3.5569 | 3.0667 | 3.0719 | 3.2333 | 3.0833 | 3.4167 | 3.2 | 3.1833
| Median |  | 3.3 | 3.8053 | 3.1959 | 6.6935 | 3 | 0.15 | 2.2 |  | 3.2 | 3 | 3.5362 | 3.05 | 3.0815 | 3.2 | 3 | 3.4 | 3.3 | 3.2
| Std Dev. |  | 0.12111 | 0.19779 | 0.12065 | 0.55782 | 0.083666 | 0.057735 | 0.05164 |  | 0.10328 | 0.05164 | 0.10296 | 0.1633 | 0.084056 | 0.05164 | 0.13292 | 0.040825 | 0.25298 | 0.075277
| Rel.Std.Dev. |  | 3.597% | 5.283% | 3.77% | 8.33% | 2.743% | 38.49% | 2.383% |  | 3.261% | 1.702% | 2.895% | 5.325% | 2.736% | 1.597% | 4.311% | 1.195% | 7.906% | 2.365%
| PDM3 |  | 5.109% | 16.87% | -0.07612% | 109.1% | -4.778% | -95.32% | -32.36% |  | -1.135% | -5.298% | 11.05% | -4.257% | -4.095% | 0.9462% | -3.737% | 6.67% | -0.09444% | -0.6148%
| Z-Score (Absolute) |  | 0.79 | 2.48 | 0.04 | 15.75 | 0.64 | 13.82 | 4.61 | 8.13 | 0.11 | 0.71 | 1.64 | 0.56 | 0.54 | 0.19 | 0.49 | 1.01 | 0.04 | 0.04
| Analytical results for Y in OREAS 254c (Certified Value 15.4 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 15.81 | 16.3542 | 15.5198 | 15.2 | 15.3 | 15 | 14.3 | 16 | 15.1 | 13.5393 | 15.7 | 14.9042 | 15.5 | 15.5 | 16.9 | 14 | 16.6 |  | 1
| 1 | 2 | 15.5 | 16.1734 | 15.3889 | 15.6 | 15.5 | 15.1 | 14.0067 | 15.8 | 14.1 | 13.6092 | 17.4 | 15.4877 | 15.3 | 15.5 | 16.2 | 14.1 | 17.1 |  | 33
| 1 | 3 | 15.44 | 16.2332 | 15.4755 | 15.7 | 15 | 15.3 | 14.3333 | 15.6 | 14.7 | 13.2394 | 15 | 15.4073 | 15.9 | 15.8 | 16.6 | 15.1 | 16.9 |  | 16
| 1 | 4 | 15.34 | 16.6677 | 15.3848 | 15.6 | 15.4 | 15.6 | 14.3567 | 15.8 | 13.9 | 13.6392 | 14.4 | 15.3522 | 16.1 | 15.1 | 16.4 | 14.8 | 16.8 |  | 15.4241
| 1 | 5 | 15.63 | 16.2895 | 15.2738 | 15.4 | 15 | 15.5 | 14.4667 | 15.6 | 14.1 | 13.3094 | 16.6 | 15.3236 | 15.8 | 22.9 | 16.2 | 14.8 | 17.3 |  | 68
| 1 | 6 | 15.52 | 16.4226 | 15.6899 | 15.7 | 15.5 | 14.8 | 14.4467 | 16.4 | 14.5 | 13.3892 | 16.3 | 14.9102 | 15.9 | 15.9 | 16.4 | 14.2 | 17.7
| Mean |  | 15.54 | 16.3567 | 15.4554 | 15.5333 | 15.2833 | 15.2167 | 14.3183 | 15.8667 | 14.4 | 13.4543 | 15.9 | 15.2309 | 15.75 | 16.7833 | 16.45 | 14.5 | 17.0667
| Median |  | 15.51 | 16.3218 | 15.4322 | 15.6 | 15.35 | 15.2 | 14.345 | 15.8 | 14.3 | 13.4642 | 16 | 15.3379 | 15.85 | 15.65 | 16.4 | 14.5 | 17
| Std Dev. |  | 0.1631 | 0.17576 | 0.14278 | 0.19664 | 0.23166 | 0.30605 | 0.16591 | 0.30111 | 0.45166 | 0.16543 | 1.0954 | 0.25688 | 0.29496 | 3.0096 | 0.26646 | 0.45607 | 0.39328
| Rel.Std.Dev. |  | 1.05% | 1.075% | 0.9238% | 1.266% | 1.516% | 2.011% | 1.159% | 1.898% | 3.137% | 1.23% | 6.89% | 1.687% | 1.873% | 17.93% | 1.62% | 3.145% | 2.304%
| PDM3 |  | 0.7515% | 6.047% | 0.2033% | 0.7082% | -0.9126% | -1.345% | -7.169% | 2.869% | -6.64% | -12.77% | 3.085% | -1.253% | 2.113% | 8.812% | 6.651% | -5.991% | 10.65%
| Z-Score (Absolute) |  | 0.01 | 1.51 | 0.14 | 0 | 0.46 | 0.58 | 2.23 | 0.61 | 2.08 | 3.82 | 0.67 | 0.56 | 0.4 | 2.3 | 1.69 | 1.9 | 2.82
| Analytical results for Yb in OREAS 254c (Certified Value 1.26 ppm) | [19 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.17 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 1.25 | 1.3799 | 1.3383 | 1.3 | 1.23 | 1.2186 | 1.2 | 1.1 | [10 empty] | 1
| 1 | 2 | 1.28 | 1.3091 | 1.3099 | 1.4 | 1.22 | 1.2485 | 1.2 | 1.1 | [10 empty] | 34
| 1 | 3 | 1.27 | 1.3083 | 1.2343 | 1.3 | 1.26 | 1.1987 | 1.2 | 1.1 | [10 empty] | 16
| 1 | 4 | 1.24 | 1.3761 | 1.3029 | 1.3 | 1.28 | 1.1687 | 1.3 | 1.1 | [10 empty] | 1.2647
| 1 | 5 | 1.3 | 1.334 | 1.2226 | 1.3 | 1.25 | 1.2089 | 1.2 | 1.2 | [10 empty] | 69
| 1 | 6 | 1.26 | 1.2562 | 1.229 | 1.4 | 1.22 | 1.2385 | 1.2 | 1.1
| Mean |  | 1.2667 | 1.3272 | 1.2728 | 1.3333 | 1.2433 | 1.2137 | 1.2167 | 1.1167
| Median |  | 1.265 | 1.3215 | 1.2686 | 1.3 | 1.24 | 1.2138 | 1.2 | 1.1
| Std Dev. |  | 0.021602 | 0.046766 | 0.049975 | 0.05164 | 0.024221 | 0.028729 | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 1.705% | 3.524% | 3.926% | 3.873% | 1.948% | 2.367% | 3.355% | 3.656%
| PDM3 |  | 0.152% | 4.941% | 0.6389% | 5.423% | -1.693% | -4.039% | -3.801% | -11.71%
| Z-Score (Absolute) |  | 0 | 1.75 | 0.18 | NA | 0.67 | 1.53 | NA | NA
Results from laboratories 1.09, 1.17 and 1.22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 254c (Certified Value 75 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 76 | 80.26 | 75.042 | 62.4647 | 75 | 75 | 70 | 64.9234 | 76 | 73 | 65.8467 | 74 | 74.7362 | 75 | 75 | 79 | 72 | 71 | 1
| 1 | 2 | 74 | 79.391 | 74.1464 | 60.7832 | 75 | 76 | 71 | 67.556 | 74 | 73 | 67.6956 | 76 | 74.8207 | 73.9 | 76 | 78 | 72 | 67 | 35
| 1 | 3 | 74 | 81.305 | 73.9337 | 61.7937 | 76 | 76 | 71 | 67.1716 | 73 | 72 | 67.1819 | 78 | 74.0111 | 74.7 | 73 | 78 | 71 | 71 | 16
| 1 | 4 | 74 | 79.914 | 74.6118 | 60.9366 | 76 | 76 | 71 | 66.8421 | 73 | 73 | 66.6656 | 75 | 72.1177 | 75.7 | 72 | 64 | 73 | 70 | 74.5566
| 1 | 5 | 73 | 80.579 | 74.9706 | 60.4353 | 75 | 76 | 71 | 66.7259 | 75 | 74 | 65.6786 | 77 | 72.1187 | 74.2 | 74 | 79 | 72 | 75 | 70
| 1 | 6 | 74 | 78.423 | 75.4455 | 62.808 | 76 | 76 | 70 | 68.4053 | 75 | 70 | 67.0041 | 76 | 73.6713 | 75.7 | 78 | 78 | 71 | 71
| Mean |  | 74.1667 | 79.9787 | 74.6917 | 61.5369 | 75.5 | 75.8333 | 70.6667 | 66.9374 | 74.3333 | 72.5 | 66.6787 | 76 | 73.5793 | 74.8667 | 74.6667 | 76 | 71.8333 | 70.8333
| Median |  | 74 | 80.087 | 74.7912 | 61.3651 | 75.5 | 76 | 71 | 67.0069 | 74.5 | 73 | 66.8348 | 76 | 73.8412 | 74.85 | 74.5 | 78 | 72 | 71
| Std Dev. |  | 0.98319 | 0.99683 | 0.57392 | 0.96793 | 0.54772 | 0.40825 | 0.5164 | 1.1578 | 1.2111 | 1.3784 | 0.78559 | 1.4142 | 1.212 | 0.75011 | 2.1602 | 5.8992 | 0.75277 | 2.5626
| Rel.Std.Dev. |  | 1.326% | 1.246% | 0.7684% | 1.573% | 0.7255% | 0.5383% | 0.7308% | 1.73% | 1.629% | 1.901% | 1.178% | 1.861% | 1.647% | 1.002% | 2.893% | 7.762% | 1.048% | 3.618%
| PDM3 |  | -0.5231% | 7.272% | 0.1811% | -17.46% | 1.265% | 1.712% | -5.217% | -10.22% | -0.2995% | -2.758% | -10.57% | 1.936% | -1.311% | 0.4158% | 0.1476% | 1.936% | -3.653% | -4.994%
| Z-Score (Absolute) |  | 0.03 | 2.21 | 0.17 | 4.9 | 0.48 | 0.61 | 1.38 | 2.82 | 0.03 | 0.67 | 2.92 | 0.67 | 0.26 | 0.24 | 0.16 | 0.67 | 0.93 | 1.32
| Analytical results for Zr in OREAS 254c (Certified Value 189 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 194.2 | 180.414 | 192.0418 | 200.4436 | 183 | 6 | 181 | 169 | 199 | 198.2 | 161.5406 | 184 | 186.4928 | 195 | 198 | 198.5 | 180.5 | 180 | 1
| 1 | 2 | 192.9 | 183.3624 | 189.9516 | 199.9341 | 185 | 6 | 181 | 170 | 192 | 195.2 | 161.4105 | 185.5 | 189.5662 | 196 | 195.5 | 193 | 183.5 | 180 | 36
| 1 | 3 | 186.9 | 182.4336 | 192.1185 | 196.5837 | 186 | 15 | 182 | 172 | 195 | 198.8 | 159.2104 | 180.5 | 185.6604 | 197 | 191.5 | 199 | 180.5 | 181 | 16
| 1 | 4 | 191.7 | 181.6344 | 190.5872 | 194.8413 | 188 | 7 | 184 | 173 | 190.5 | 199.6 | 160.9032 | 180 | 189.2345 | 200 | 188 | 184 | 193.5 | 175 | 188.7761
| 1 | 5 | 192.3 | 179.982 | 193.9315 | 196.4418 | 189 | 15 | 183 | 172 | 184.5 | 199.2 | 158.1854 | 207 | 184.0011 | 195 | 191 | 195 | 195.5 | 181 | 71
| 1 | 6 | 189.9 | 182.1204 | 195.6507 | 203.1242 | 188 | 8 | 182 | 174 | 193.5 | 192.2 | 161.4702 | 191.5 | 188.3508 | 199 | 198.5 | 195.5 | 186 | 186
| Mean |  | 191.3167 | 181.6578 | 192.3802 | 198.5615 | 186.5 | 9.5 | 182.1667 | 171.6667 | 192.4167 | 197.2 | 160.4534 | 188.0833 | 187.2176 | 197 | 193.75 | 194.1667 | 186.5833 | 180.5
| Median |  | 192 | 181.8774 | 192.0801 | 198.2589 | 187 | 7.5 | 182 | 172 | 192.75 | 198.5 | 161.1568 | 184.75 | 187.4218 | 196.5 | 193.5 | 195.25 | 184.75 | 180.5
| Std Dev. |  | 2.5864 | 1.271 | 2.1165 | 3.1141 | 2.2583 | 4.3243 | 1.169 | 1.8619 | 4.852 | 2.9093 | 1.4158 | 10.1559 | 2.1984 | 2.0976 | 4.2279 | 5.465 | 6.4994 | 3.5071
| Rel.Std.Dev. |  | 1.352% | 0.6997% | 1.1% | 1.568% | 1.211% | 45.52% | 0.6417% | 1.085% | 2.522% | 1.475% | 0.8824% | 5.4% | 1.174% | 1.065% | 2.182% | 2.815% | 3.483% | 1.943%
| PDM3 |  | 1.346% | -3.771% | 1.909% | 5.184% | -1.206% | -94.97% | -3.501% | -9.063% | 1.929% | 4.462% | -15% | -0.367% | -0.8255% | 4.356% | 2.635% | 2.856% | -1.162% | -4.384%
| Z-Score (Absolute) |  | 0.41 | 0.67 | 0.53 | 1.22 | 0.13 | 19.87 | 0.61 | 1.78 | 0.53 | 1.07 | 3.03 | 0.05 | 0.05 | 1.04 | 0.68 | 0.73 | 0.12 | 0.8
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C925: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C943: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C961: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 254c (Certified Value 0.928 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.93 | 0.90848 | < 10 | 0.9 | 0.927 | 0.89 | 0.90193 | 0.93 | 0.9 | 1.089 | 0.9 | 0.84186 | 1 | 0.95 | 0.93 | 0.95 |  |  |  | 1
| 1 | 2 | 0.93 | 0.91541 | < 10 | 0.91 | 0.971 | 0.92 | 0.89537 | 0.94 | 1 | 1.0804 | 0.9 | 0.79581 | 0.9 | 0.93 | 0.97 | 0.94 |  |  |  | 20
| 1 | 3 | 0.92 | 0.905 | < 10 | 0.91 | 0.997 | 0.91 | 0.91163 | 0.92 | 0.9 | 1.092 | 0.9 | 0.80474 | 1 | 0.94 | 0.91 | 0.98 |  |  |  | 16
| 1 | 4 | 0.93 | 0.93337 | < 10 | 0.87 | 0.97 | 0.93 | 0.8776 | 0.92 | 0.9 | 1.0872 | 0.99 | 0.8477 | 0.9 | 0.9 | 0.98 | 0.97 |  |  |  | 0.92813
| 1 | 5 | 0.93 | 0.90478 | < 10 | 0.92 | 0.946 | 0.91 | 0.8785 | 0.9 | 1 | 1.087 | 0.93 | 0.79816 | 1 | 0.95 | 0.93 | 0.96 |  |  |  | 73
| 1 | 6 | 0.95 | 0.89963 | < 10 | 0.87 | 0.924 | 0.91 | 0.9182 | 0.92 | 0.9 | 1.0704 | 0.95 | 0.83363 | 0.9 | 0.94 | 1 | 0.96
| Mean |  | 0.93167 | 0.91111 |  | 0.89667 | 0.95583 | 0.91167 | 0.89721 | 0.92167 | 0.93333 | 1.0843 | 0.92833 | 0.82032 | 0.95 | 0.935 | 0.95333 | 0.96
| Median |  | 0.93 | 0.90674 |  | 0.905 | 0.958 | 0.91 | 0.89865 | 0.92 | 0.9 | 1.0871 | 0.915 | 0.81918 | 0.95 | 0.94 | 0.95 | 0.96
| Std Dev. |  | 0.0098319 | 0.012082 |  | 0.021602 | 0.02852 | 0.013292 | 0.016787 | 0.013292 | 0.05164 | 0.0078166 | 0.03656 | 0.023348 | 0.054772 | 0.018708 | 0.035024 | 0.014142
| Rel.Std.Dev. |  | 1.055% | 1.326% |  | 2.409% | 2.984% | 1.458% | 1.871% | 1.442% | 5.533% | 0.7209% | 3.938% | 2.846% | 5.766% | 2.001% | 3.674% | 1.473%
| PDM3 |  | 0.3805% | -1.834% |  | -3.39% | 2.984% | -1.774% | -3.332% | -0.6969% | 0.5601% | 16.83% | 0.02138% | -11.62% | 2.356% | 0.7397% | 2.715% | 3.433%
| Z-Score (Absolute) |  | 0.05 | 0.52 | 111.64 | 0.91 | 0.71 | 0.5 | 0.9 | 0.23 | NA | 4.23 | 0.05 | 3.01 | NA | 0.14 | 0.64 | 0.82
Results from laboratories 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 254c (Certified Value 1.69 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.6748 | 1.9648 | 1.656 | 1.48 | 1.93 | 1.78 | 1.45 | 1.78 | 1.62 | 1.58 | 1.54 | 1.8422 | 1.93 | 1.74 | 1.65 | 1.63 | 1.502 |  |  | 1
| 1 | 2 | 1.7254 | 1.9909 | 1.697 | 1.32 | 1.95 | 1.82 | 1.5 | 1.76 | 1.56 | 1.62 | 1.52 | 1.8062 | 1.93 | 1.66 | 1.69 | 1.6 | 1.514 |  |  | #N/A
| 1 | 3 | 1.7398 | 1.9128 | 1.6688 | 1.53 | 1.92 | 1.83 | 1.57 | 1.8 | 1.59 | 1.61 | 1.53 | 1.7825 | 1.91 | 1.74 | 1.56 | 1.58 | 1.496 |  |  | 16
| 1 | 4 | 1.7179 | 1.9024 | 1.7162 | 1.44 | 1.87 | 1.77 | 1.59 | 1.77 | 1.6 | 1.56 | 1.61 | 1.7825 | 1.93 | 1.66 | 1.73 | 1.59 | 1.455 |  |  | 1.6879
| 1 | 5 | 1.7059 | 1.975 | 1.6863 | 1.33 | 1.89 | 1.83 | 1.56 | 1.79 | 1.56 | 1.67 | 1.57 | 1.8581 | 1.91 | 1.72 | 1.63 | 1.58 | 1.414 |  |  | 74
| 1 | 6 | 1.6522 | 1.9221 | 1.6985 | 1.4 | 1.91 | 1.77 | 1.47 | 1.73 | 1.6 | 1.6 | 1.63 | 1.8273 | 1.93 | 1.7 | 1.72 | 1.59 | 1.464
| Mean |  | 1.7027 | 1.9447 | 1.6871 | 1.4167 | 1.9117 | 1.8 | 1.5233 | 1.7717 | 1.5883 | 1.6067 | 1.5667 | 1.8165 | 1.9233 | 1.7033 | 1.6633 | 1.595 | 1.4742
| Median |  | 1.7119 | 1.9435 | 1.6917 | 1.42 | 1.915 | 1.8 | 1.53 | 1.775 | 1.595 | 1.605 | 1.555 | 1.8168 | 1.93 | 1.71 | 1.67 | 1.59 | 1.48
| Std Dev. |  | 0.03305 | 0.036806 | 0.021801 | 0.083106 | 0.028577 | 0.029665 | 0.05785 | 0.024833 | 0.024014 | 0.037771 | 0.045019 | 0.031411 | 0.010328 | 0.036697 | 0.06377 | 0.018708 | 0.037226
| Rel.Std.Dev. |  | 1.941% | 1.893% | 1.292% | 5.866% | 1.495% | 1.648% | 3.798% | 1.402% | 1.512% | 2.351% | 2.874% | 1.729% | 0.537% | 2.154% | 3.834% | 1.173% | 2.525%
| PDM3 |  | 0.872% | 15.21% | -0.04773% | -16.07% | 13.25% | 6.638% | -9.752% | 4.96% | -5.902% | -4.815% | -7.185% | 7.615% | 13.95% | 0.9115% | -1.458% | -5.507% | -12.67%
| Z-Score (Absolute) |  | 0.09 | 1.54 | 0 | 1.62 | 1.34 | 0.67 | 0.98 | 0.51 | 0.59 | 0.48 | 0.72 | 0.77 | 1.41 | 0.1 | 0.14 | 0.55 | 1.27
| Analytical results for As in OREAS 254c (Certified Value 9.25 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 9 | 9.4387 | < 100 | 11 | 10 | 9.4 | 10.3 | 8.84 | 8.7 | 9 | 8.54 | 8.6 | 8.6584 | 9.8 | 8.8 | 9.1 | 8.2 | 11 |  | 1
| 1 | 2 | 9 | 9.5194 | < 100 | 12 | 10 | 9.5 | 10.3 | 9.3633 | 8.5 | 9.4 | 8.505 | 8.5 | 8.9137 | 9.8 | 8.5 | 9.2 | 8.5 | 11 |  | 21
| 1 | 3 | 9 | 9.7237 | < 100 | 11 | 10 | 9.5 | 10.6 | 9.2567 | 8.3 | 8.7 | 8.517 | 8.5 | 8.6581 | 10.3 | 8.7 | 9.3 | 8.9 | 11 |  | 16
| 1 | 4 | 9 | 9.8745 | < 100 | 11 | 10 | 9.2 | 10.6 | 8.36 | 8.4 | 8.5 | 8.5437 | 9.1 | 8.6484 | 10.3 | 8.3 | 10.5 | 8.5 | 11 |  | 9.2498
| 1 | 5 | 9 | 9.7053 | < 100 | 36 | 9 | 9.4 | 10.7 | 9.1167 | 8.7 | 9.5 | 8.505 | 8.8 | 8.6383 | 10.2 | 8.8 | 9.3 | 8.4 | 10 |  | 75
| 1 | 6 | 9 | 9.6628 | < 100 | 14 | 9 | 9.1 | 10.3 | 8.2333 | 8.2 | 8.5 | 8.554 | 9 | 9.2269 | 10 | 8.6 | 10.2 | 8.5 | 11
| Mean |  | 9 | 9.6541 |  | 15.8333 | 9.6667 | 9.35 | 10.4667 | 8.8617 | 8.4667 | 8.9333 | 8.5274 | 8.75 | 8.7906 | 10.0667 | 8.6167 | 9.6 | 8.5 | 10.8333
| Median |  | 9 | 9.684 |  | 11.5 | 10 | 9.4 | 10.45 | 8.9783 | 8.45 | 8.85 | 8.5285 | 8.7 | 8.6582 | 10.1 | 8.65 | 9.3 | 8.5 | 11
| Std Dev. |  | 0 | 0.15535 |  | 9.9482 | 0.5164 | 0.16432 | 0.18619 | 0.47321 | 0.20656 | 0.44121 | 0.021183 | 0.25884 | 0.23833 | 0.23381 | 0.19408 | 0.5933 | 0.22804 | 0.40825
| Rel.Std.Dev. |  | 0% | 1.609% |  | 62.83% | 5.342% | 1.757% | 1.779% | 5.34% | 2.44% | 4.939% | 0.2484% | 2.958% | 2.711% | 2.323% | 2.252% | 6.18% | 2.683% | 3.768%
| PDM3 |  | -2.7% | 4.371% |  | 71.18% | 4.507% | 1.083% | 13.16% | -4.196% | -8.466% | -3.421% | -7.809% | -5.403% | -4.964% | 8.831% | -6.845% | 3.786% | -8.106% | 17.12%
| Z-Score (Absolute) |  | 0.22 | 0.62 | 52.44 | 8.55 | 0.63 | 0.22 | 1.66 | 0.4 | 0.91 | 0.31 | 0.83 | 0.55 | 0.49 | 1.15 | 0.72 | 0.55 | 0.87 | 2.13
| Analytical results for B in OREAS 254c (Certified Value < 10 ppm) | [20 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [9 empty] | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 1
|  |  | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 1
| 1 | 1 | < 10 | 1.1 | 4 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 1
| 1 | 2 | < 10 | 1.6 | 3 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | #N/A
| 1 | 3 | < 10 | 13.3 | 4 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 16
| 1 | 4 | < 10 | 4.6 | 4 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | < 10
| 1 | 5 | < 10 | 14.4 | 3 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 76
| 1 | 6 | < 10 | 5.1 | 4 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10
| Mean |  |  | 6.6833 | 3.6667
| Median |  |  | 4.85 | 4
| Std Dev. |  |  | 5.7825 | 0.5164
| Rel.Std.Dev. |  |  | 86.52% | 14.08%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 254c (Certified Value 95 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 93 | 103.248 | 93.432 | 98.843 | 104 | 92 | 87 | 93 | 94.0167 | 100 | 91 | 90 | 95.1208 | 100 | 100 | 90 | 84 |  |  | 1
| 1 | 2 | 95 | 108.7776 | 96.046 | 101.827 | 98 | 92 | 91.8 | 95 | 92.6633 | 90 | 90 | 90 | 92.5518 | 100 | 100 | 90 | 83 |  |  | 22
| 1 | 3 | 94 | 105.7968 | 95.1336 | 95.515 | 108 | 91 | 93.5 | 98 | 95.9233 | 90 | 86 | 90 | 97.1513 | 100 | 90 | 90 | 85 |  |  | 16
| 1 | 4 | 97 | 102.0168 | 95.9299 | 100.299 | 104 | 88 | 90 | 97 | 90.2467 | 90 | 86 | 100 | 92.2653 | 100 | 100 | 90 | 84 |  |  | 94.5444
| 1 | 5 | 93 | 107.3628 | 95.4548 | 105.528 | 100 | 89 | 90.2 | 97 | 95.24 | 90 | 92 | 90 | 97.446 | 100 | 90 | 90 | 82 |  |  | 77
| 1 | 6 | 94 | 104.5224 | 96.6042 | 103.611 | 105 | 93 | 86.8 | 94 | 89.54 | 90 | 82 | 100 | 94.7827 | 100 | 100 | 90 | 82
| Mean |  | 94.3333 | 105.2874 | 95.4334 | 100.9372 | 103.1667 | 90.8333 | 89.8833 | 95.6667 | 92.9383 | 91.6667 | 87.8333 | 93.3333 | 94.8863 | 100 | 96.6667 | 90 | 83.3333
| Median |  | 94 | 105.1596 | 95.6923 | 101.063 | 104 | 91.5 | 90.1 | 96 | 93.34 | 90 | 88 | 90 | 94.9518 | 100 | 100 | 90 | 83.5
| Std Dev. |  | 1.5055 | 2.5373 | 1.1035 | 3.5555 | 3.6009 | 1.9408 | 2.6339 | 1.9664 | 2.6169 | 4.0825 | 3.8166 | 5.164 | 2.194 | 0 | 5.164 | 0 | 1.2111
| Rel.Std.Dev. |  | 1.596% | 2.41% | 1.156% | 3.523% | 3.49% | 2.137% | 2.93% | 2.055% | 2.816% | 4.454% | 4.345% | 5.533% | 2.312% | 0% | 5.342% | 0% | 1.453%
| PDM3 |  | -0.2232% | 11.36% | 0.9403% | 6.762% | 9.12% | -3.925% | -4.93% | 1.187% | -1.699% | -3.044% | -7.098% | -1.281% | 0.3617% | 5.77% | 2.245% | -4.807% | -11.86%
| Z-Score (Absolute) |  | 0.04 | 1.69 | 0.13 | 1 | 1.36 | 0.6 | 0.75 | 0.17 | 0.27 | NA | 1.07 | NA | 0.04 | NA | NA | NA | 1.79
Results from laboratories 1.12, 1.15, 1.19, 1.20 and 1.21 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 254c (Certified Value 1.08 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 1.09 | 1.0783 | < 10 | 1.2 | 1 | 1.1 | 0.96 | < 5 | 1.03 | 0.99 | 1.0409 | 1.2 | 1.11 | 1.07 | 1.02 | 1.31 |  |  |  | 1
| 1 | 2 | 1.06 | 1.1174 | < 10 | 1.3 | 1 | 1.2 | 0.98 | < 5 | 1 | 1 | 0.9774 | 1.1 | 1.07 | 1.09 | 1 | 1.22 |  |  |  | 23
| 1 | 3 | 1.09 | 1.087 | < 10 | 1.2 | 1.1 | 1.2 | 1.02 | < 5 | 1.02 | 0.99 | 0.99183 | 1.1 | 1.1 | 1.02 | 1 | 1.29 |  |  |  | 16
| 1 | 4 | 1.11 | 1.0755 | < 10 | 1.2 | 1 | 1 | 1.01 | < 5 | 1.03 | 1.05 | 0.9659 | 1.1 | 1.07 | 1.08 | 1 | 1.21 |  |  |  | 1.0797
| 1 | 5 | 1.08 | 1.1067 | < 10 | 1.2 | 1 | 1.1 | 0.99 | < 5 | 1 | 1.02 | 1.1056 | 1.1 | 1.1 | 1.06 | 1 | 1.22 |  |  |  | 78
| 1 | 6 | 1.07 | 1.0953 | < 10 | 1.2 | 1 | 1.1 | 0.95 | < 5 | 1.02 | 1.06 | 1.0498 | 1.2 | 1.09 | 1.11 | 1 | 1.24
| Mean |  | 1.0833 | 1.0934 |  | 1.2167 | 1.0167 | 1.1167 | 0.985 |  | 1.0167 | 1.0183 | 1.0219 | 1.1333 | 1.09 | 1.0717 | 1.0033 | 1.2483
| Median |  | 1.085 | 1.0912 |  | 1.2 | 1 | 1.1 | 0.985 |  | 1.02 | 1.01 | 1.0164 | 1.1 | 1.095 | 1.075 | 1 | 1.23
| Std Dev. |  | 0.017512 | 0.016419 |  | 0.040825 | 0.040825 | 0.075277 | 0.027386 |  | 0.013663 | 0.030605 | 0.053212 | 0.05164 | 0.016733 | 0.030605 | 0.008165 | 0.041673
| Rel.Std.Dev. |  | 1.616% | 1.502% |  | 3.355% | 4.016% | 6.741% | 2.78% |  | 1.344% | 3.005% | 5.207% | 4.556% | 1.535% | 2.856% | 0.8138% | 3.338%
| PDM3 |  | 0.34% | 1.27% |  | 12.69% | -5.835% | 3.427% | -8.768% |  | -5.835% | -5.68% | -5.35% | 4.971% | 0.9575% | -0.7406% | -7.07% | 15.62%
| Z-Score (Absolute) |  | 0.03 | 0.07 | 38.15 | 1.27 | 0.68 | 0.29 | 0.99 | 13.78 | 0.68 | 0.67 | 0.63 | 0.45 | 0.03 | 0.15 | 0.81 | 1.58
| Analytical results for Bi in OREAS 254c (Certified Value 0.31 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.32 | 0.30225 | < 10 | < 0.3 | 0.3 | 0.3 | 0.25 | < 2 | 0.3 | 0.3 | 0.307 | 0.28 | 0.29968 | 0.35 | 0.32 | 0.32 | 0.28 | 0.5 |  | 1
| 1 | 2 | 0.33 | 0.35347 | < 10 | < 0.3 | 0.31 | 0.32 | 0.26 | < 2 | 0.29 | 0.3 | 0.3052 | 0.29 | 0.35315 | 0.35 | 0.31 | 0.34 | 0.28 | 0.57 |  | 24
| 1 | 3 | 0.32 | 0.31566 | < 10 | < 0.3 | 0.3 | 0.33 | 0.28 | < 2 | 0.29 | 0.3 | 0.305 | 0.27 | 0.32103 | 0.33 | 0.31 | 0.38 | 0.27 | 0.57 |  | 16
| 1 | 4 | 0.32 | 0.31504 | < 10 | < 0.3 | 0.32 | 0.31 | 0.28 | < 2 | 0.3 | 0.3 | 0.307 | 0.32 | 0.3066 | 0.35 | 0.31 | 0.33 | 0.28 | 0.53 |  | 0.3104
| 1 | 5 | 0.32 | 0.35876 | < 10 | < 0.3 | 0.31 | 0.31 | 0.27 | < 2 | 0.29 | 0.3 | 0.306 | 0.29 | 0.34908 | 0.35 | 0.32 | 0.33 | 0.29 | 0.51 |  | 79
| 1 | 6 | 0.32 | 0.32957 | < 10 | < 0.3 | 0.32 | 0.3 | 0.26 | < 2 | 0.29 | 0.3 | 0.306 | 0.31 | 0.35069 | 0.35 | 0.32 | 0.33 | 0.29 | 0.53
| Mean |  | 0.32167 | 0.32912 |  |  | 0.31 | 0.31167 | 0.26667 |  | 0.29333 | 0.3 | 0.30603 | 0.29333 | 0.33004 | 0.34667 | 0.315 | 0.33833 | 0.28167 | 0.535
| Median |  | 0.32 | 0.32261 |  |  | 0.31 | 0.31 | 0.265 |  | 0.29 | 0.3 | 0.306 | 0.29 | 0.33505 | 0.35 | 0.315 | 0.33 | 0.28 | 0.53
| Std Dev. |  | 0.0040825 | 0.022686 |  |  | 0.0089443 | 0.01169 | 0.012111 |  | 0.005164 | 0 | 0.00085245 | 0.018619 | 0.023982 | 0.008165 | 0.0054772 | 0.02137 | 0.0075277 | 0.029496
| Rel.Std.Dev. |  | 1.269% | 6.893% |  |  | 2.885% | 3.751% | 4.541% |  | 1.76% | 0% | 0.2785% | 6.347% | 7.267% | 2.355% | 1.739% | 6.316% | 2.673% | 5.513%
| PDM3 |  | 3.63% | 6.032% |  |  | -0.1287% | 0.4083% | -14.09% |  | -5.498% | -3.35% | -1.407% | -5.498% | 6.326% | 11.68% | 1.482% | 8.999% | -9.257% | 72.36%
| Z-Score (Absolute) |  | 0.21 | 0.44 | 145.81 | 5.14 | 0.16 | 0.1 | 1.5 | 21.32 | 0.67 | NA | 0.28 | 0.67 | 0.47 | 0.99 | 0 | 0.73 | 1.04 | 6.85
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 254c (Certified Value 0.473 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 0.49 | 0.57118 | 0.48084 | 0.39494 | 1.11 | 0.51 | 0.62 | 0.43 | 0.5 | 0.48 | 0.4 | 0.39629 | 0.44 | 0.50376 | 0.57 | 0.52 | 0.48 | 0.46 | 0.4 | 1
| 1 | 2 | 0.5 | 0.57371 | 0.48769 | 0.40197 | 1.06 | 0.51 | 0.64 | 0.43 | 0.5 | 0.46 | 0.41 | 0.3953 | 0.44 | 0.49031 | 0.55 | 0.49 | 0.49 | 0.45 | 0.404 | #N/A
| 1 | 3 | 0.51 | 0.57303 | 0.4816 | 0.38009 | 1.11 | 0.51 | 0.62 | 0.43 | 0.52 | 0.47 | 0.46 | 0.41718 | 0.44 | 0.48619 | 0.55 | 0.52 | 0.46 | 0.45 | 0.399 | 16
| 1 | 4 | 0.5 | 0.57209 | 0.4926 | 0.39731 | 1.08 | 0.5 | 0.62 | 0.43 | 0.48 | 0.48 | 0.44 | 0.398 | 0.47 | 0.47661 | 0.56 | 0.5 | 0.49 | 0.45 | 0.389 | 0.4727
| 1 | 5 | 0.5 | 0.57235 | 0.48725 | 0.39897 | 1.05 | 0.5 | 0.63 | 0.43 | 0.51 | 0.46 | 0.48 | 0.39119 | 0.46 | 0.51492 | 0.56 | 0.51 | 0.48 | 0.45 | 0.379 | 80
| 1 | 6 | 0.48 | 0.5713 | 0.48898 | 0.39958 | 1.26 | 0.51 | 0.57 | 0.42 | 0.48 | 0.47 | 0.46 | 0.3936 | 0.47 | 0.48991 | 0.57 | 0.5 | 0.5 | 0.45 | 0.389
| Mean |  | 0.49667 | 0.57227 | 0.48649 | 0.39548 | 1.1117 | 0.50667 | 0.61667 | 0.42833 | 0.49833 | 0.47 | 0.44167 | 0.39859 | 0.45333 | 0.49362 | 0.56 | 0.50667 | 0.48333 | 0.45167 | 0.39333
| Median |  | 0.5 | 0.57222 | 0.48747 | 0.39814 | 1.095 | 0.51 | 0.62 | 0.43 | 0.5 | 0.47 | 0.45 | 0.3958 | 0.45 | 0.49011 | 0.56 | 0.505 | 0.485 | 0.45 | 0.394
| Std Dev. |  | 0.010328 | 0.00098085 | 0.004502 | 0.0078945 | 0.07679 | 0.005164 | 0.024221 | 0.0040825 | 0.016021 | 0.0089443 | 0.031252 | 0.0094002 | 0.015055 | 0.013607 | 0.0089443 | 0.012111 | 0.013663 | 0.0040825 | 0.0093095
| Rel.Std.Dev. |  | 2.079% | 0.1714% | 0.9254% | 1.996% | 6.908% | 1.019% | 3.928% | 0.9531% | 3.215% | 1.903% | 7.076% | 2.358% | 3.321% | 2.757% | 1.597% | 2.39% | 2.827% | 0.9039% | 2.367%
| PDM3 |  | 5.071% | 21.07% | 2.919% | -16.34% | 135.2% | 7.187% | 30.46% | -9.385% | 5.424% | -0.5701% | -6.564% | -15.68% | -4.096% | 4.426% | 18.47% | 7.187% | 2.251% | -4.449% | -16.79%
| Z-Score (Absolute) |  | 0.2 | 1.66 | 0 | 1.76 | 12.1 | 0.39 | 2.52 | 1.13 | 0.23 | 0.32 | 0.87 | 1.7 | 0.64 | 0.14 | 1.42 | 0.39 | 0.06 | 0.67 | 1.8
| Analytical results for Cd in OREAS 254c (Certified Value 0.055 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.04 | 0.11 | < 0.5 | 0.05 | 0.1 | 0.05 | 0.072368 | 0.09 | 0.05 | 0.04 | 0.05 |  |  |  | 1
| 1 | 2 | 0.04 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.03 | 0.12 | < 0.5 | 0.05 | < 0.1 | 0.05 | 0.072012 | 0.09 | 0.04 | 0.05 | 0.04 |  |  |  | 25
| 1 | 3 | 0.05 | < 0.2 | < 10 | < 0.3 | 0.1 | 0.07 | 0.12 | < 0.5 | 0.06 | < 0.1 | 0.05 | 0.050307 | 0.08 | 0.05 | 0.05 | 0.04 |  |  |  | 16
| 1 | 4 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.11 | 0.07 | 0.12 | < 0.5 | 0.05 | < 0.1 | 0.06 | 0.05404 | 0.08 | 0.05 | 0.04 | 0.04 |  |  |  | 0.054611
| 1 | 5 | 0.05 | < 0.2 | < 10 | < 0.3 | 0.1 | 0.09 | 0.12 | < 0.5 | 0.04 | < 0.1 | 0.05 | 0.057453 | 0.07 | 0.05 | 0.04 | 0.04 |  |  |  | 81
| 1 | 6 | 0.05 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.07 | 0.1 | < 0.5 | 0.04 | < 0.1 | 0.06 | 0.052809 | 0.08 | 0.05 | 0.05 | 0.04
| Mean |  | 0.051667 |  |  |  | 0.093333 | 0.061667 | 0.115 |  | 0.048333 | 0.1 | 0.053333 | 0.059831 | 0.081667 | 0.048333 | 0.045 | 0.041667
| Median |  | 0.05 |  |  |  | 0.095 | 0.07 | 0.12 |  | 0.05 | 0.1 | 0.05 | 0.055747 | 0.08 | 0.05 | 0.045 | 0.04
| Std Dev. |  | 0.0075277 |  |  |  | 0.012111 | 0.022286 | 0.0083666 |  | 0.0075277 |  | 0.005164 | 0.0098464 | 0.0075277 | 0.0040825 | 0.0054772 | 0.0040825
| Rel.Std.Dev. |  | 14.57% |  |  |  | 12.98% | 36.14% | 7.275% |  | 15.57% |  | 9.682% | 16.46% | 9.218% | 8.447% | 12.17% | 9.798%
| PDM3 |  | -5.391% |  |  |  | 70.91% | 12.92% | 110.6% |  | -11.5% | 83.11% | -2.339% | 9.56% | 49.54% | -11.5% | -17.6% | -23.7%
| Z-Score (Absolute) |  | 0.35 | 1.52 | 191.24 | 3.46 | 1.26 | 0.04 | 2.1 | 7.33 | 0.48 | 0.09 | 0.29 | 0.04 | 0.81 | 0.48 | 0.61 | 0.74
| Analytical results for Ce in OREAS 254c (Certified Value 39.7 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 38.428 | 37.3993 | 40.6 | 40.9 | 40 | 40.4 | 38.9 | 40.7891 | 35.3 | 42.9 | 42.2 | 39.6 | 44.79 | [6 empty] | 1
| 1 | 2 | 39.386 | 37.8117 | 39.7 | 41.6 | 39 | 38.5 | 39 | 40.9656 | 35.8 | 40.3 | 43.1 | 38.6 | 43.13 | [6 empty] | 26
| 1 | 3 | 38.947 | 37.4312 | 40.6 | 41.8 | 41 | 38.7 | 38.6 | 39.8333 | 36.7 | 42 | 41.6 | 38.6 | 42.84 | [6 empty] | 16
| 1 | 4 | 40.474 | 38.509 | 39.1 | 40.1 | 39 | 39.7 | 41.9 | 39.9811 | 36.8 | 39.6 | 44.4 | 38.3 | 43 | [6 empty] | 39.7114
| 1 | 5 | 38.212 | 37.527 | 39 | 40.4 | 40 | 38.9 | 41.3 | 42.0589 | 37.3 | 42.4 | 46.2 | 37.6 | 43.22 | [6 empty] | 82
| 1 | 6 | 39.081 | 38.0017 | 40.1 | 40.1 | 39 | 39.2 | 40.5 | 41.6538 | 36.3 | 41.4 | 45.8 | 38.1 | 42.42
| Mean |  | 39.088 | 37.78 | 39.85 | 40.8167 | 39.6667 | 39.2333 | 40.0333 | 40.8803 | 36.3667 | 41.4333 | 43.8833 | 38.4667 | 43.2333
| Median |  | 39.014 | 37.6694 | 39.9 | 40.65 | 39.5 | 39.05 | 39.75 | 40.8773 | 36.5 | 41.7 | 43.75 | 38.45 | 43.065
| Std Dev. |  | 0.80408 | 0.42711 | 0.7064 | 0.74677 | 0.8165 | 0.70899 | 1.3938 | 0.88404 | 0.72572 | 1.2691 | 1.8957 | 0.66833 | 0.81299
| Rel.Std.Dev. |  | 2.057% | 1.131% | 1.773% | 1.83% | 2.058% | 1.807% | 3.482% | 2.163% | 1.996% | 3.063% | 4.32% | 1.737% | 1.88%
| PDM3 |  | -1.57% | -4.864% | 0.349% | 2.783% | -0.1127% | -1.204% | 0.8107% | 2.943% | -8.423% | 4.336% | 10.51% | -3.134% | 8.869%
| Z-Score (Absolute) |  | 0.5 | 1.35 | 0 | 0.63 | 0.12 | 0.4 | 0.12 | 0.67 | 2.28 | 1.04 | 2.64 | 0.91 | 2.21
| Analytical results for Co in OREAS 254c (Certified Value 16.7 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 16 | 18.5868 | 16.6151 | 20.466 | 16 | 16.9 | 17.7 | 15.5 | 17.3133 | 15.9 | 14.9 | 16.3 | 16.7692 | 17.3 | 17.1 | 15.9 | 15.7 | 19.2 |  | 1
| 1 | 2 | 15.9 | 19.062 | 16.4366 | 19.603 | 15 | 16.6 | 18 | 15.8 | 17.1033 | 15.4 | 14.6 | 16.3 | 16.4737 | 17.27 | 16.5 | 16.3 | 15.2 | 18.5 |  | 27
| 1 | 3 | 16 | 18.1224 | 16.4727 | 18.059 | 16 | 16.9 | 18.2 | 16.3 | 17.73 | 15.6 | 15.6 | 16.4 | 16.4068 | 17.83 | 17.2 | 15.7 | 15.7 | 19.1 |  | 16
| 1 | 4 | 16.4 | 17.766 | 16.8011 | 17.939 | 28 | 16.8 | 17.5 | 16.8 | 16.8433 | 15.6 | 15.5 | 17.3 | 16.4093 | 17.59 | 16 | 16.8 | 15.7 | 18.1 |  | 16.7363
| 1 | 5 | 16 | 18.7164 | 16.5758 | 20.233 | 30 | 16.8 | 17.7 | 16.4 | 17.65 | 15.5 | 16 | 16.7 | 16.8396 | 17.77 | 16.8 | 17 | 15.3 | 18.2 |  | 83
| 1 | 6 | 16.1 | 17.9064 | 16.6751 | 18.35 | 13 | 16.6 | 17.5 | 16 | 16.58 | 15.4 | 14.9 | 17.1 | 16.7484 | 17.42 | 16.7 | 16.9 | 15.8 | 18.1
| Mean |  | 16.0667 | 18.36 | 16.5961 | 19.1083 | 19.6667 | 16.7667 | 17.7667 | 16.1333 | 17.2033 | 15.5667 | 15.25 | 16.6833 | 16.6078 | 17.53 | 16.7167 | 16.4333 | 15.5667 | 18.5333
| Median |  | 16 | 18.3546 | 16.5955 | 18.9765 | 16 | 16.8 | 17.7 | 16.15 | 17.2083 | 15.55 | 15.2 | 16.55 | 16.6111 | 17.505 | 16.75 | 16.55 | 15.7 | 18.35
| Std Dev. |  | 0.17512 | 0.50721 | 0.13396 | 1.131 | 7.3394 | 0.13663 | 0.28048 | 0.46332 | 0.45103 | 0.18619 | 0.53198 | 0.43089 | 0.19867 | 0.23824 | 0.43551 | 0.55015 | 0.25033 | 0.50067
| Rel.Std.Dev. |  | 1.09% | 2.763% | 0.8072% | 5.919% | 37.32% | 0.8149% | 1.579% | 2.872% | 2.622% | 1.196% | 3.488% | 2.583% | 1.196% | 1.359% | 2.605% | 3.348% | 1.608% | 2.701%
| PDM3 |  | -4.001% | 9.702% | -0.8378% | 14.17% | 17.51% | 0.1815% | 6.157% | -3.603% | 2.791% | -6.989% | -8.881% | -0.3164% | -0.7675% | 4.742% | -0.1172% | -1.81% | -6.989% | 10.74%
| Z-Score (Absolute) |  | 0.67 | 1.83 | 0.1 | 2.65 | NA | 0.09 | 1.18 | 0.6 | 0.57 | 1.22 | 1.57 | 0 | 0.08 | 0.93 | 0.04 | 0.27 | 1.22 | 2.02
Results from laboratory 1.06 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cr in OREAS 254c (Certified Value 60 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 0
| 1 | 1 | 64 | 69.3144 | 59.5067 | 45.822 | 49 | 76 | 65 | 55 | 62.2207 | 60 | 54 | 55 | 68.2229 | 70 | 63 | 61 | 58 | 57 |  | 1
| 1 | 2 | 66 | 69.5628 | 60.2483 | 46.823 | 44 | 76 | 67 | 55 | 62.2066 | 57 | 55 | 55 | 66.263 | 70 | 60 | 61 | 56 | 58 |  | 28
| 1 | 3 | 66 | 66.4416 | 59.4949 | 44.357 | 50 | 75 | 68 | 57 | 63.3555 | 59 | 56 | 55 | 67.6657 | 72 | 63 | 58 | 56 | 58 |  | 16
| 1 | 4 | 65 | 65.9556 | 60.7188 | 46.473 | 47 | 73 | 65 | 56 | 61.3357 | 59 | 55 | 57 | 66.0555 | 72 | 61 | 61 | 56 | 55 |  | 60.4829
| 1 | 5 | 65 | 68.31 | 59.8333 | 49.64 | 44 | 74 | 66 | 55 | 63.2878 | 58 | 60 | 57 | 67.5939 | 72 | 63 | 61 | 56 | 55 |  | 84
| 1 | 6 | 63 | 66.96 | 60.5851 | 48.306 | 45 | 74 | 64 | 54 | 60.3889 | 59 | 56 | 59 | 69.0041 | 71 | 62 | 64 | 56 | 56
| Mean |  | 64.8333 | 67.7574 | 60.0645 | 46.9035 | 46.5 | 74.6667 | 65.8333 | 55.3333 | 62.1325 | 58.6667 | 56 | 56.3333 | 67.4675 | 71.1667 | 62 | 61 | 56.3333 | 56.5
| Median |  | 65 | 67.635 | 60.0408 | 46.648 | 46 | 74.5 | 65.5 | 55 | 62.2137 | 59 | 55.5 | 56 | 67.6298 | 71.5 | 62.5 | 61 | 56 | 56.5
| Std Dev. |  | 1.169 | 1.5233 | 0.53327 | 1.8602 | 2.5884 | 1.2111 | 1.472 | 1.0328 | 1.1418 | 1.0328 | 2.0976 | 1.633 | 1.1342 | 0.98319 | 1.2649 | 1.8974 | 0.8165 | 1.3784
| Rel.Std.Dev. |  | 1.803% | 2.248% | 0.8878% | 3.966% | 5.567% | 1.622% | 2.236% | 1.866% | 1.838% | 1.76% | 3.746% | 2.899% | 1.681% | 1.382% | 2.04% | 3.11% | 1.449% | 2.44%
| PDM3 |  | 7.193% | 12.03% | -0.6917% | -22.45% | -23.12% | 23.45% | 8.846% | -8.514% | 2.727% | -3.003% | -7.412% | -6.861% | 11.55% | 17.66% | 2.508% | 0.855% | -6.861% | -6.585%
| Z-Score (Absolute) |  | 0.66 | 1.1 | 0.07 | 2.08 | 2.14 | 2.16 | 0.81 | 0.79 | 0.24 | 0.28 | 0.69 | 0.64 | 1.06 | 1.62 | 0.22 | 0.07 | 0.64 | 0.62
| Analytical results for Cs in OREAS 254c (Certified Value 0.92 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.93 | 0.88493 | 1.12 | 0.9 | 0.93 | 0.86 | 0.8 | 1.0334 | 1.23 | 0.9 | 0.93 | 0.81 | 0.98 | [6 empty] | 1
| 1 | 2 | 0.98 | 0.88442 | 1.15 | 0.88 | 0.97 | 0.82 | 0.78 | 0.95667 | 1.22 | 0.85 | 0.94 | 0.81 | 1 | [6 empty] | 29
| 1 | 3 | 0.96 | 0.87288 | 1.16 | 0.95 | 0.96 | 0.84 | 0.78 | 0.91351 | 1.24 | 0.88 | 0.88 | 0.81 | 1 | [6 empty] | 16
| 1 | 4 | 1 | 0.91156 | 1.11 | 0.88 | 1 | 0.85 | 0.81 | 0.92726 | 1.23 | 0.83 | 0.98 | 0.83 | 1.04 | [6 empty] | 0.92483
| 1 | 5 | 0.96 | 0.91039 | 1.11 | 0.87 | 0.99 | 0.84 | 0.82 | 1.051 | 1.24 | 0.88 | 0.94 | 0.81 | 0.99 | [6 empty] | 85
| 1 | 6 | 0.98 | 0.91861 | 1.13 | 0.92 | 0.92 | 0.84 | 0.84 | 1.0096 | 1.22 | 0.87 | 0.95 | 0.81 | 1
| Mean |  | 0.96833 | 0.89713 | 1.13 | 0.9 | 0.96167 | 0.84167 | 0.805 | 0.98189 | 1.23 | 0.86833 | 0.93667 | 0.81333 | 1.0017
| Median |  | 0.97 | 0.89766 | 1.125 | 0.89 | 0.965 | 0.84 | 0.805 | 0.98311 | 1.23 | 0.875 | 0.94 | 0.81 | 1
| Std Dev. |  | 0.024014 | 0.018675 | 0.020976 | 0.030332 | 0.031885 | 0.013292 | 0.023452 | 0.057416 | 0.0089443 | 0.024833 | 0.03266 | 0.008165 | 0.020412
| Rel.Std.Dev. |  | 2.48% | 2.082% | 1.856% | 3.37% | 3.316% | 1.579% | 2.913% | 5.847% | 0.7272% | 2.86% | 3.487% | 1.004% | 2.038%
| PDM3 |  | 4.703% | -2.996% | 22.18% | -2.685% | 3.983% | -8.993% | -12.96% | 6.169% | 33% | -6.109% | 1.279% | -12.06% | 8.308%
| Z-Score (Absolute) |  | 0.33 | 0.41 | 2.01 | 0.38 | 0.26 | 0.99 | 1.37 | 0.47 | 3.04 | 0.71 | 0 | 1.28 | 0.67
| Analytical results for Cu in OREAS 254c (Certified Value 26.9 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 2
| 1 | 1 | 26.9 | 23.5116 | 26.7629 | 23.316 | 30 | 26 | 29.5 | 27 | 29.0467 | 25.1 | 29.7 | 24.5542 | 24 | 24.419 | 29.1 | 29.8 | 29.8 | 26.1 | 23.4 | 1
| 1 | 2 | 27.5 | 24.624 | 27.4489 | 24.596 | 30 | 27 | 29.5 | 26 | 28.7667 | 24.8 | 30.2 | 24.8957 | 24.5 | 25.5138 | 29.2 | 28.9 | 30.1 | 26.5 | 23.4 | 30
| 1 | 3 | 27.2 | 23.2632 | 26.8868 | 22.591 | 30 | 26 | 28.6 | 27 | 29.82 | 24.3 | 28 | 25.1452 | 24.1 | 24.6565 | 30 | 29.1 | 28.5 | 26.7 | 23.3 | 16
| 1 | 4 | 27.5 | 23.1832 | 27.4927 | 23.581 | 29 | 27 | 29.1 | 27 | 28.05 | 24.8 | 26.8 | 24.4756 | 26.5 | 24.0903 | 29.5 | 27.9 | 30.5 | 27.6 | 22.9 | 26.8533
| 1 | 5 | 27 | 23.7708 | 27.3155 | 24.478 | 29 | 28 | 28.7 | 27 | 29.72 | 24.4 | 29.9 | 24.3685 | 25.5 | 25.1913 | 30 | 29.6 | 31 | 27.5 | 23.6 | 86
| 1 | 6 | 27.7 | 22.5288 | 27.2918 | 23.559 | 28 | 29 | 28.6 | 27 | 27.82 | 24.7 | 28.1 | 24.8757 | 25.9 | 26.4446 | 29.6 | 29.2 | 31.1 | 26.9 | 22.3
| Mean |  | 27.3 | 23.4803 | 27.1998 | 23.6868 | 29.3333 | 27.1667 | 29 | 26.8333 | 28.8706 | 24.6833 | 28.7833 | 24.7191 | 25.0833 | 25.0526 | 29.5667 | 29.0833 | 30.1667 | 26.8833 | 23.15
| Median |  | 27.35 | 23.3874 | 27.3037 | 23.57 | 29.5 | 27 | 28.9 | 27 | 28.9067 | 24.75 | 28.9 | 24.715 | 25 | 24.9239 | 29.55 | 29.15 | 30.3 | 26.8 | 23.35
| Std Dev. |  | 0.31623 | 0.69774 | 0.30285 | 0.75079 | 0.8165 | 1.169 | 0.42895 | 0.40825 | 0.82967 | 0.29269 | 1.3497 | 0.29888 | 1.0323 | 0.85487 | 0.38297 | 0.66758 | 0.95847 | 0.58109 | 0.47645
| Rel.Std.Dev. |  | 1.158% | 2.972% | 1.113% | 3.17% | 2.784% | 4.303% | 1.479% | 1.521% | 2.874% | 1.186% | 4.689% | 1.209% | 4.116% | 3.412% | 1.295% | 2.295% | 3.177% | 2.162% | 2.058%
| PDM3 |  | 1.664% | -12.56% | 1.29% | -11.79% | 9.236% | 1.167% | 7.994% | -0.07431% | 7.512% | -8.081% | 7.187% | -7.947% | -6.591% | -6.706% | 10.1% | 8.305% | 12.34% | 0.1119% | -13.79%
| Z-Score (Absolute) |  | 0.04 | 1.19 | 0.01 | 1.13 | 0.7 | 0 | 0.59 | 0.11 | 0.55 | 0.8 | 0.52 | 0.79 | 0.67 | 0.68 | 0.78 | 0.62 | 0.97 | 0.09 | 1.3
| Analytical results for Dy in OREAS 254c (Indicative Value 2.16 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 2.196 | 2.0237 | 2.2 | [16 empty] | 1
| 1 | 2 | 2.221 | 2.1165 | 2.2 | [16 empty] | 4
| 1 | 3 | 2.185 | 2.0302 | 2.4 | [16 empty] | 16
| 1 | 4 | 2.221 | 2.1965 | 2.2 | [16 empty] | 2.162
| 1 | 5 | 2.141 | 2.0716 | 2.1 | [16 empty] | 10
| 1 | 6 | 2.132 | 2.0818 | 2.2
| Mean |  | 2.1827 | 2.0867 | 2.2167
| Median |  | 2.1905 | 2.0767 | 2.2
| Std Dev. |  | 0.038537 | 0.0638 | 0.098319
| Rel.Std.Dev. |  | 1.766% | 3.057% | 4.435%
| PDM3 |  | 0.9552% | -3.483% | 2.528%
| Z-Score (Absolute) |  | 0 | 1.9 | 0.67
| Analytical results for Er in OREAS 254c (Indicative Value 0.85 ppm) | [20 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.923 | 0.79806 | 0.8 | [16 empty] | 1
| 1 | 2 | 0.908 | 0.86646 | 0.8 | [16 empty] | 5
| 1 | 3 | 0.894 | 0.81009 | 0.8 | [16 empty] | 16
| 1 | 4 | 0.913 | 0.86507 | 0.8 | [16 empty] | 0.84583
| 1 | 5 | 0.881 | 0.81356 | 0.8 | [16 empty] | 11
| 1 | 6 | 0.917 | 0.83577 | 0.8
| Mean |  | 0.906 | 0.8315 | 0.8
| Median |  | 0.9105 | 0.82467 | 0.8
| Std Dev. |  | 0.015697 | 0.029209 | 1.2162e-16
| Rel.Std.Dev. |  | 1.733% | 3.513% | 1.52e-14%
| PDM3 |  | 7.113% | -1.694% | -5.419%
| Z-Score (Absolute) |  | 1.59 | 0 | 0.67
| Analytical results for Eu in OREAS 254c (Indicative Value 0.98 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 1.054 | 0.8922 | 1 | [16 empty] | 1
| 1 | 2 | 1.078 | 0.91459 | 1 | [16 empty] | 6
| 1 | 3 | 1.022 | 0.89067 | 1 | [16 empty] | 16
| 1 | 4 | 1.074 | 0.97595 | 0.9 | [16 empty] | 0.97834
| 1 | 5 | 1.028 | 0.90419 | 1 | [16 empty] | 12
| 1 | 6 | 1.038 | 0.93862 | 0.9
| Mean |  | 1.049 | 0.91937 | 0.96667
| Median |  | 1.046 | 0.90939 | 1
| Std Dev. |  | 0.023588 | 0.032822 | 0.05164
| Rel.Std.Dev. |  | 2.249% | 3.57% | 5.342%
| PDM3 |  | 7.222% | -6.028% | -1.194%
| Z-Score (Absolute) |  | 1.17 | 0.67 | 0
| Analytical results for Fe in OREAS 254c (Certified Value 2.98 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 3.02 | 3.2927 | 2.9981 | 2.6357 | 3.03 | 3.32 | 3.49 | 2.64 | 2.99 | 2.96 | 2.75 | 2.7346 | 2.81 | 3.2539 | 3.36 | 3.13 | 2.9 | 2.89 |  | 1
| 1 | 2 | 3.11 | 3.4387 | 3.0391 | 2.6696 | 2.8 | 3.35 | 3.48 | 2.64 | 2.98 | 2.86 | 2.83 | 2.8036 | 2.84 | 3.2724 | 3.31 | 3 | 2.93 | 2.83 |  | #N/A
| 1 | 3 | 3.11 | 3.2978 | 2.9912 | 2.5731 | 3.13 | 3.32 | 3.37 | 2.64 | 3.08 | 2.9 | 2.91 | 2.765 | 2.79 | 3.1538 | 3.3 | 3.11 | 2.73 | 2.83 |  | 16
| 1 | 4 | 3.04 | 3.2482 | 3.0536 | 2.6344 | 2.95 | 3.23 | 3.41 | 2.64 | 2.92 | 2.93 | 2.8 | 2.7771 | 2.95 | 3.158 | 3.37 | 3.01 | 3 | 2.82 |  | 2.9783
| 1 | 5 | 3.06 | 3.2738 | 2.9977 | 2.6332 | 2.8 | 3.23 | 3.38 | 2.64 | 3.05 | 2.87 | 3.05 | 2.7122 | 2.86 | 3.2372 | 3.35 | 3.08 | 2.85 | 2.81 |  | 87
| 1 | 6 | 2.99 | 3.29 | 3.0455 | 2.6578 | 2.89 | 3.27 | 3.4 | 2.63 | 2.97 | 2.93 | 2.91 | 2.7963 | 2.97 | 3.2638 | 3.35 | 3.06 | 3 | 2.84
| Mean |  | 3.055 | 3.3069 | 3.0209 | 2.634 | 2.9333 | 3.2867 | 3.4217 | 2.6383 | 2.9983 | 2.9083 | 2.875 | 2.7648 | 2.87 | 3.2232 | 3.34 | 3.065 | 2.9017 | 2.8367
| Median |  | 3.05 | 3.2914 | 3.0186 | 2.6351 | 2.92 | 3.295 | 3.405 | 2.64 | 2.985 | 2.915 | 2.87 | 2.7711 | 2.85 | 3.2455 | 3.35 | 3.07 | 2.915 | 2.83
| Std Dev. |  | 0.048477 | 0.067062 | 0.028094 | 0.033296 | 0.13095 | 0.050859 | 0.051153 | 0.0040825 | 0.057764 | 0.038687 | 0.10616 | 0.035554 | 0.074027 | 0.053421 | 0.028284 | 0.05244 | 0.10226 | 0.028048
| Rel.Std.Dev. |  | 1.587% | 2.028% | 0.93% | 1.264% | 4.464% | 1.547% | 1.495% | 0.1547% | 1.927% | 1.33% | 3.693% | 1.286% | 2.579% | 1.657% | 0.8468% | 1.711% | 3.524% | 0.9888%
| PDM3 |  | 2.574% | 11.03% | 1.428% | -11.56% | -1.511% | 10.35% | 14.89% | -11.42% | 0.6715% | -2.35% | -3.469% | -7.17% | -3.637% | 8.221% | 12.14% | 2.91% | -2.574% | -4.757%
| Z-Score (Absolute) |  | 0.53 | 2.01 | 0.33 | 1.96 | 0.19 | 1.89 | 2.69 | 1.93 | 0.19 | 0.34 | 0.54 | 1.19 | 0.57 | 1.52 | 2.21 | 0.59 | 0.38 | 0.76
| Analytical results for Ga in OREAS 254c (Certified Value 6.05 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 5.8 | 5.5983 | 7 | 6.85 | 5.99 | 5.9709 | 5.71 | 5 | 5.09 | 6.2704 | 6.57 | 5.97 | 6.2 | 5.19 | [5 empty] | 1
| 1 | 2 | 5.88 | 5.771 | 6.9 | 6.88 | 6.12 | 5.8986 | 5.63 | 5 | 5 | 5.9095 | 6.65 | 5.63 | 6.36 | 5.24 | [5 empty] | 34
| 1 | 3 | 5.86 | 5.7104 | 6.7 | 7.06 | 6.35 | 6.1887 | 5.54 | 5 | 5.11 | 6.0197 | 7.01 | 5.85 | 6.15 | 5.26 | [5 empty] | 16
| 1 | 4 | 5.98 | 5.9072 | 6.7 | 6.84 | 6.51 | 5.7475 | 5.71 | 5 | 5.65 | 5.8985 | 6.82 | 5.53 | 6.53 | 5.22 | [5 empty] | 6.0462
| 1 | 5 | 5.72 | 5.7629 | 6.7 | 6.83 | 6.35 | 6.1339 | 5.5 | 5 | 5.47 | 6.2764 | 6.73 | 5.85 | 6.73 | 5.14 | [5 empty] | 88
| 1 | 6 | 5.66 | 5.9423 | 7 | 6.79 | 6.02 | 5.669 | 5.64 | 5 | 5.53 | 6.0789 | 6.58 | 5.77 | 6.8 | 5.22
| Mean |  | 5.8167 | 5.782 | 6.8333 | 6.875 | 6.2233 | 5.9348 | 5.6217 | 5 | 5.3083 | 6.0756 | 6.7267 | 5.7667 | 6.4617 | 5.2117
| Median |  | 5.83 | 5.767 | 6.8 | 6.845 | 6.235 | 5.9347 | 5.635 | 5 | 5.29 | 6.0493 | 6.69 | 5.81 | 6.445 | 5.22
| Std Dev. |  | 0.11553 | 0.12707 | 0.15055 | 0.095237 | 0.21011 | 0.20615 | 0.086583 | 0 | 0.27353 | 0.16755 | 0.16789 | 0.1612 | 0.27096 | 0.042151
| Rel.Std.Dev. |  | 1.986% | 2.198% | 2.203% | 1.385% | 3.376% | 3.474% | 1.54% | 0% | 5.153% | 2.758% | 2.496% | 2.795% | 4.193% | 0.8088%
| PDM3 |  | -3.796% | -4.369% | 13.02% | 13.71% | 2.93% | -1.843% | -7.021% | -17.3% | -12.2% | 0.4859% | 11.25% | -4.623% | 6.872% | -13.8%
| Z-Score (Absolute) |  | 0.25 | 0.33 | 1.94 | 2.02 | 0.62 | 0 | 0.67 | NA | 1.35 | 0.3 | 1.71 | 0.36 | 1.13 | 1.56
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 254c (Indicative Value 3.24 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 3.201 | 3.1395 | 3.4 | [16 empty] | 1
| 1 | 2 | 3.276 | 3.3261 | 3.3 | [16 empty] | 7
| 1 | 3 | 3.261 | 3.2589 | 3.2 | [16 empty] | 16
| 1 | 4 | 3.269 | 3.2597 | 3.2 | [16 empty] | 3.2425
| 1 | 5 | 3.179 | 3.2405 | 3.2 | [16 empty] | 13
| 1 | 6 | 3.157 | 3.2968 | 3.2
| Mean |  | 3.2238 | 3.2536 | 3.25
| Median |  | 3.231 | 3.2593 | 3.2
| Std Dev. |  | 0.051266 | 0.063856 | 0.083666
| Rel.Std.Dev. |  | 1.59% | 1.963% | 2.574%
| PDM3 |  | -0.5748% | 0.3426% | 0.2322%
| Z-Score (Absolute) |  | 4.93 | 0.67 | 0
| Analytical results for Ge in OREAS 254c (Certified Value 0.11 ppm) | [20 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.09 | 1.10 | 1.12 | 1.15 | 1.19 | 1.20 | 1.21 | 1.22 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [10 empty] | 2
|  |  | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | [10 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.13 | 0.14 | 0.09 | 0.09 | 0.11 | 0.08 | 0.1 | [10 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.12 | 0.13 | 0.11 | 0.09 | 0.16 | 0.08 | 0.1 | [10 empty] | 1
| 1 | 3 | < 0.2 | < 0.1 | 0.12 | 0.11 | 0.09 | 0.09 | 0.11 | 0.08 | 0.1 | [10 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.13 | 0.11 | 0.11 | 0.09 | 0.16 | 0.08 | 0.1 | [10 empty] | 0.10881
| 1 | 5 | < 0.2 | < 0.1 | 0.13 | 0.12 | 0.1 | 0.08 | 0.14 | 0.08 | 0.1 | [10 empty] | 89
| 1 | 6 | < 0.2 | < 0.1 | 0.13 | 0.13 | 0.11 | 0.09 | 0.17 | 0.08 | 0.1
| Mean |  |  |  | 0.12667 | 0.12333 | 0.10167 | 0.088333 | 0.14167 | 0.08 | 0.1
| Median |  |  |  | 0.13 | 0.125 | 0.105 | 0.09 | 0.15 | 0.08 | 0.1
| Std Dev. |  |  |  | 0.005164 | 0.012111 | 0.0098319 | 0.0040825 | 0.026394 | 0 | 1.5202e-17
| Rel.Std.Dev. |  |  |  | 4.077% | 9.819% | 9.671% | 4.622% | 18.63% | 0% | 1.52e-14%
| PDM3 |  |  |  | 16.41% | 13.35% | -6.565% | -18.82% | 30.2% | -26.48% | -8.096%
| Z-Score (Absolute) |  | 0 | 1.69 | 0.9 | 0.79 | 0.06 | 0.39 | 1.4 | 0.67 | 0
| Analytical results for Hf in OREAS 254c (Indicative Value 0.66 ppm) | [20 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 0.14 | 0.74832 | 1.24 | < 0.1 | 0.61 | 0.47 | 0.52 | 1.0388 | 0.93 | 0.37 | 0.57 | 0.31 | 1.14 | [6 empty] | 1
| 1 | 2 | 0.16 | 0.70456 | 1.23 | < 0.1 | 0.65 | 0.53 | 0.63 | 0.79974 | 0.93 | 0.49 | 0.53 | 0.32 | 1.11 | [6 empty] | 8
| 1 | 3 | 0.19 | 0.71058 | 1.15 | < 0.1 | 0.68 | 0.56 | 0.51 | 0.94943 | 0.89 | 0.36 | 0.5 | 0.41 | 1.14 | [6 empty] | 16
| 1 | 4 | 0.16 | 0.93094 | 1.12 | < 0.1 | 0.63 | 0.56 | 0.4 | 0.70415 | 0.84 | 0.38 | 0.64 | 0.33 | 1.1 | [6 empty] | 0.66385
| 1 | 5 | 0.14 | 0.73885 | 1.17 | < 0.1 | 0.71 | 0.55 | 0.47 | 1.042 | 0.97 | 0.42 | 0.5 | 0.28 | 1.06 | [6 empty] | 14
| 1 | 6 | 0.16 | 0.94302 | 1.15 | < 0.1 | 0.61 | 0.52 | 0.46 | 0.79902 | 0.96 | 0.39 | 0.53 | 0.33 | 1.08
| Mean |  | 0.15833 | 0.79604 | 1.1767 |  | 0.64833 | 0.53167 | 0.49833 | 0.88885 | 0.92 | 0.40167 | 0.545 | 0.33 | 1.105
| Median |  | 0.16 | 0.74358 | 1.16 |  | 0.64 | 0.54 | 0.49 | 0.87458 | 0.93 | 0.385 | 0.53 | 0.325 | 1.105
| Std Dev. |  | 0.018348 | 0.11047 | 0.048028 |  | 0.040208 | 0.034303 | 0.077309 | 0.14122 | 0.048166 | 0.047924 | 0.053198 | 0.043359 | 0.032094
| Rel.Std.Dev. |  | 11.59% | 13.88% | 4.082% |  | 6.202% | 6.452% | 15.51% | 15.89% | 5.235% | 11.93% | 9.761% | 13.14% | 2.904%
| PDM3 |  | -76.15% | 19.91% | 77.25% |  | -2.338% | -19.91% | -24.93% | 33.89% | 38.59% | -39.49% | -17.9% | -50.29% | 66.45%
| Z-Score (Absolute) |  | 1.04 | 0.67 | 1.7 | 1.33 | 0.28 | 0.04 | 0.13 | 0.92 | 1.01 | 0.38 | 0 | 0.58 | 1.5
| Analytical results for Hg in OREAS 254c (Certified Value 0.035 ppm) | [20 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [5 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | < 0.1 | < 1 | 0.02 | < 2 | 0.04 | < 0.01 | 0.029 | 0.02 | < 0.05 | 1 | 0.04 | 0.02 | 0.04 | 0.26 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 1 | 0.07 | < 2 | 0.04 | 0.06 | 0.03 | 0.03 | < 0.05 | 1 | 0.04 | 0.03 | 0.05 | 0.24 | [5 empty] | 2
| 1 | 3 | < 0.1 | < 1 | 0.04 | < 2 | 0.05 | 0.04 | 0.03 | 0.04 | < 0.05 | 1 | 0.04 | < 0.01 | 0.05 | 0.27 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 1 | 0.02 | < 2 | 0.04 | 0.03 | 0.033 | 0.04 | < 0.05 | 1 | 0.04 | 0.02 | 0.05 | 0.26 | [5 empty] | 0.034979
| 1 | 5 | < 0.1 | < 1 | 0.02 | < 2 | 0.04 | 0.05 | 0.028 | 0.03 | < 0.05 | 1 | 0.03 | < 0.01 | 0.04 | 0.25 | [5 empty] | 90
| 1 | 6 | < 0.1 | < 1 | 0.04 | < 2 | 0.03 | 0.02 | 0.031 | 0.04 | < 0.05 | 1 | 0.04 | < 0.01 | 0.05 | 0.26
| Mean |  |  |  | 0.035 |  | 0.04 | 0.04 | 0.030167 | 0.033333 |  | 1 | 0.038333 | 0.023333 | 0.046667 | 0.25667
| Median |  |  |  | 0.03 |  | 0.04 | 0.04 | 0.03 | 0.035 |  | 1 | 0.04 | 0.02 | 0.05 | 0.26
| Std Dev. |  |  |  | 0.019748 |  | 0.0063246 | 0.015811 | 0.0017224 | 0.008165 |  | 0 | 0.0040825 | 0.0057735 | 0.005164 | 0.010328
| Rel.Std.Dev. |  |  |  | 56.42% |  | 15.81% | 39.53% | 5.71% | 24.49% |  | 0% | 10.65% | 24.74% | 11.07% | 4.024%
| PDM3 |  |  |  | 0.05956% |  | 14.35% | 14.35% | -13.76% | -4.705% |  | 2759% | 9.589% | -33.29% | 33.41% | 633.8%
| Z-Score (Absolute) |  | 0.94 | 37.36 | 0.27 | 77.82 | 0.13 | 0.34 | 0.66 | 0.4 | 1.08 | NA | 0 | 1.96 | 0.67 | 17.67
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 254c (Indicative Value 0.36 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.348 | 0.33676 | 0.4 | [16 empty] | 1
| 1 | 2 | 0.363 | 0.32672 | 0.4 | [16 empty] | 9
| 1 | 3 | 0.36 | 0.33231 | 0.4 | [16 empty] | 16
| 1 | 4 | 0.365 | 0.33613 | 0.4 | [16 empty] | 0.35808
| 1 | 5 | 0.34 | 0.3409 | 0.3 | [16 empty] | 15
| 1 | 6 | 0.357 | 0.33956 | 0.4
| Mean |  | 0.3555 | 0.3354 | 0.38333
| Median |  | 0.3585 | 0.33644 | 0.4
| Std Dev. |  | 0.0096488 | 0.0051915 | 0.040825
| Rel.Std.Dev. |  | 2.714% | 1.548% | 10.65%
| PDM3 |  | -0.7197% | -6.334% | 7.053%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.93
| Analytical results for In in OREAS 254c (Certified Value 0.022 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.02 | < 0.05 | 0.03 | < 5 | 0.023 | 0.024 | < 0.05 | 0.1 | 0.022 | 0.022 | 0.022 | 0.02 | [7 empty] | 1
| 1 | 2 | 0.02 | < 0.05 | 0.03 | < 5 | 0.02 | 0.023 | < 0.05 | 0.1 | 0.021 | 0.025 | 0.021 | 0.02 | [7 empty] | 22
| 1 | 3 | 0.02 | < 0.05 | 0.03 | < 5 | 0.022 | 0.017 | < 0.05 | 0.1 | 0.022 | 0.023 | 0.021 | 0.023 | [7 empty] | 16
| 1 | 4 | 0.02 | < 0.05 | 0.03 | < 5 | 0.021 | 0.024 | < 0.05 | 0.1 | 0.021 | 0.026 | 0.022 | 0.018 | [7 empty] | 0.021595
| 1 | 5 | 0.02 | < 0.05 | 0.03 | < 5 | 0.023 | 0.019 | < 0.05 | 0.1 | 0.022 | 0.025 | 0.02 | 0.021 | [7 empty] | 91
| 1 | 6 | 0.02 | < 0.05 | 0.03 | < 5 | 0.021 | 0.027 | < 0.05 | 0.1 | 0.02 | 0.024 | 0.022 | 0.02
| Mean |  | 0.02 |  | 0.03 |  | 0.021667 | 0.022333 |  | 0.1 | 0.021333 | 0.024167 | 0.021333 | 0.020333
| Median |  | 0.02 |  | 0.03 |  | 0.0215 | 0.0235 |  | 0.1 | 0.0215 | 0.0245 | 0.0215 | 0.02
| Std Dev. |  | 0 |  | 0 |  | 0.0012111 | 0.0036697 |  | 1.5202e-17 | 0.0008165 | 0.001472 | 0.0008165 | 0.001633
| Rel.Std.Dev. |  | 0% |  | 0% |  | 5.59% | 16.43% |  | 1.52e-14% | 3.827% | 6.091% | 3.827% | 8.031%
| PDM3 |  | -7.387% |  | 38.92% |  | 0.3308% | 3.418% |  | 363.1% | -1.213% | 11.91% | -1.213% | -5.843%
| Z-Score (Absolute) |  | 0.79 | 0.9 | 2.58 | 835.36 | 0.22 | 0 | 0.9 | NA | 0.34 | 0.62 | 0.34 | 0.67
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 254c (Certified Value 0.295 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.288 | 0.29593 | 0.2533 | 0.37 | 0.321 | 0.31 | 0.23 | 0.31 | 0.3 | 0.27 | 0.28 | 0.32366 | 0.36 | 0.32 | 0.29 | 0.28 | 0.24 |  |  | 1
| 1 | 2 | 0.2981 | 0.30126 | 0.25645 | 0.34 | 0.322 | 0.3 | 0.23 | 0.31 | 0.29 | 0.28 | 0.28 | 0.31502 | 0.35 | 0.31 | 0.29 | 0.28 | 0.24 |  |  | #N/A
| 1 | 3 | 0.3 | 0.29404 | 0.24291 | 0.38 | 0.326 | 0.29 | 0.24 | 0.32 | 0.3 | 0.28 | 0.27 | 0.31553 | 0.35 | 0.32 | 0.28 | 0.27 | 0.236 |  |  | 16
| 1 | 4 | 0.2953 | 0.30315 | 0.25483 | 0.36 | 0.323 | 0.3 | 0.24 | 0.3 | 0.3 | 0.26 | 0.29 | 0.30905 | 0.35 | 0.3 | 0.29 | 0.27 | 0.231 |  |  | 0.29518
| 1 | 5 | 0.2934 | 0.29775 | 0.25607 | 0.35 | 0.324 | 0.29 | 0.24 | 0.32 | 0.29 | 0.29 | 0.29 | 0.3296 | 0.36 | 0.31 | 0.29 | 0.27 | 0.226 |  |  | 92
| 1 | 6 | 0.2884 | 0.30427 | 0.2573 | 0.4 | 0.324 | 0.3 | 0.22 | 0.3 | 0.3 | 0.27 | 0.3 | 0.31776 | 0.35 | 0.31 | 0.3 | 0.27 | 0.232
| Mean |  | 0.29387 | 0.2994 | 0.25347 | 0.36667 | 0.32333 | 0.29833 | 0.23333 | 0.31 | 0.29667 | 0.275 | 0.285 | 0.31844 | 0.35333 | 0.31167 | 0.29 | 0.27333 | 0.23417
| Median |  | 0.29435 | 0.29951 | 0.25545 | 0.365 | 0.3235 | 0.3 | 0.235 | 0.31 | 0.3 | 0.275 | 0.285 | 0.31665 | 0.35 | 0.31 | 0.29 | 0.27 | 0.234
| Std Dev. |  | 0.0049419 | 0.0041191 | 0.005362 | 0.021602 | 0.0017512 | 0.0075277 | 0.008165 | 0.0089443 | 0.005164 | 0.010488 | 0.010488 | 0.0072171 | 0.005164 | 0.0075277 | 0.0063246 | 0.005164 | 0.0055287
| Rel.Std.Dev. |  | 1.682% | 1.376% | 2.115% | 5.892% | 0.5416% | 2.523% | 3.499% | 2.885% | 1.741% | 3.814% | 3.68% | 2.266% | 1.462% | 2.415% | 2.181% | 1.889% | 2.361%
| PDM3 |  | -0.4462% | 1.429% | -14.13% | 24.22% | 9.536% | 1.067% | -20.95% | 5.019% | 0.5023% | -6.838% | -3.45% | 7.877% | 19.7% | 5.584% | -1.756% | -7.402% | -20.67%
| Z-Score (Absolute) |  | 0.09 | 0.09 | 1.34 | 2.18 | 0.83 | 0.05 | 1.97 | 0.41 | 0 | 0.67 | 0.36 | 0.68 | 1.76 | 0.47 | 0.21 | 0.73 | 1.95
| Analytical results for La in OREAS 254c (Certified Value 22.7 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 21.077 | 21.2371 | 21.9 | 23.3 | 20.5 | 21.9767 | 22.8 | 22 | 22.5 | 22.9813 | 17 | 24.4 | 23.9 | 22.6 | 22.9 |  |  |  |  | 1
| 1 | 2 | 21.57 | 21.7008 | 21.8 | 24 | 20.9 | 21.8433 | 22.6 | 21 | 22.9 | 22.8853 | 17.1 | 23 | 25 | 22.1 | 23.9 |  |  |  |  | 23
| 1 | 3 | 22.017 | 21.269 | 22 | 24.3 | 21.7 | 22.5733 | 22.5 | 21 | 23 | 22.5266 | 18 | 24 | 23.9 | 22.2 | 23.8 |  |  |  |  | 16
| 1 | 4 | 21.892 | 21.4432 | 22 | 23.4 | 22.1 | 21.43 | 22.9 | 21 | 24.6 | 22.5708 | 17.4 | 22.8 | 25.8 | 22 | 24.4 |  |  |  |  | 22.6649
| 1 | 5 | 21.292 | 21.7675 | 21.9 | 23.1 | 21.7 | 22.48 | 22.8 | 22 | 23.9 | 23.6279 | 17.8 | 24.1 | 26.4 | 21.8 | 24.5 |  |  |  |  | 93
| 1 | 6 | 21.102 | 21.7216 | 22.4 | 23 | 20.7 | 21.08 | 22.5 | 22 | 24.4 | 23.5849 | 17.6 | 23.7 | 26.2 | 22 | 23.7
| Mean |  | 21.4917 | 21.5232 | 22 | 23.5167 | 21.2667 | 21.8972 | 22.6833 | 21.5 | 23.55 | 23.0295 | 17.4833 | 23.6667 | 25.2 | 22.1167 | 23.8667
| Median |  | 21.431 | 21.572 | 21.95 | 23.35 | 21.3 | 21.91 | 22.7 | 21.5 | 23.45 | 22.9333 | 17.5 | 23.85 | 25.4 | 22.05 | 23.85
| Std Dev. |  | 0.40151 | 0.23811 | 0.20976 | 0.51929 | 0.65013 | 0.58174 | 0.17224 | 0.54772 | 0.86891 | 0.48019 | 0.392 | 0.6377 | 1.1153 | 0.27142 | 0.57504
| Rel.Std.Dev. |  | 1.868% | 1.106% | 0.9535% | 2.208% | 3.057% | 2.657% | 0.7593% | 2.548% | 3.69% | 2.085% | 2.242% | 2.695% | 4.426% | 1.227% | 2.409%
| PDM3 |  | -5.176% | -5.037% | -2.933% | 3.758% | -6.169% | -3.387% | 0.08145% | -5.14% | 3.905% | 1.609% | -22.86% | 4.42% | 11.19% | -2.419% | 5.302%
| Z-Score (Absolute) |  | 0.5 | 0.47 | 0.09 | 1.11 | 0.67 | 0.17 | 0.45 | 0.49 | 1.14 | 0.72 | 3.68 | 1.23 | 2.45 | 0 | 1.39
| Analytical results for Li in OREAS 254c (Certified Value 8.73 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 8.5 | 8.6625 | 8 | 9.7 | 9.1 | 7.5 | 8.8318 | 8.3 | 8 | 5.3843 | 10 | 9.2 | 8.4 | 8.3 | 10 |  |  |  |  | 1
| 1 | 2 | 8.6 | 8.8665 | 8 | 9.5 | 9.3 | 7.9 | 8.7616 | 8.1 | 7.9 | 4.5959 | 10 | 8.8 | 8.5 | 8.1 | 11 |  |  |  |  | #N/A
| 1 | 3 | 8.5 | 8.7915 | 9 | 9.7 | 9.3 | 8.1 | 9.1291 | 8.3 | 7.9 | 4.4341 | 10.3 | 9.2 | 8 | 8 | 10 |  |  |  |  | 16
| 1 | 4 | 8.8 | 9.0728 | 8 | 9.3 | 9.2 | 7.9 | 8.351 | 8.2 | 8.2 | 5.3466 | 10.1 | 8.8 | 8.2 | 8.1 | 10 |  |  |  |  | 8.728
| 1 | 5 | 8.5 | 8.8128 | 8 | 9.5 | 9.3 | 7.9 | 9.03 | 7.8 | 8.2 | 4.5139 | 10.5 | 9.2 | 8.3 | 8.1 | 9 |  |  |  |  | 94
| 1 | 6 | 8.6 | 9.0049 | 9 | 9.7 | 9.2 | 7.3 | 8.3816 | 8.2 | 8.5 | 5.293 | 10 | 8.9 | 8.7 | 8.1 | 10
| Mean |  | 8.5833 | 8.8685 | 8.3333 | 9.5667 | 9.2333 | 7.7667 | 8.7475 | 8.15 | 8.1167 | 4.928 | 10.15 | 9.0167 | 8.35 | 8.1167 | 10
| Median |  | 8.55 | 8.8397 | 8 | 9.6 | 9.25 | 7.9 | 8.7967 | 8.2 | 8.1 | 4.9445 | 10.05 | 9.05 | 8.35 | 8.1 | 10
| Std Dev. |  | 0.1169 | 0.14954 | 0.5164 | 0.1633 | 0.08165 | 0.30111 | 0.32367 | 0.18708 | 0.23166 | 0.4566 | 0.20736 | 0.20412 | 0.2429 | 0.098319 | 0.63246
| Rel.Std.Dev. |  | 1.362% | 1.686% | 6.197% | 1.707% | 0.8843% | 3.877% | 3.7% | 2.295% | 2.854% | 9.265% | 2.043% | 2.264% | 2.909% | 1.211% | 6.325%
| PDM3 |  | -1.658% | 1.61% | -4.522% | 9.609% | 5.79% | -11.01% | 0.2235% | -6.622% | -7.004% | -43.54% | 16.29% | 3.307% | -4.331% | -7.004% | 14.57%
| Z-Score (Absolute) |  | 0 | 0.41 | NA | 1.42 | 0.94 | 1.18 | 0.24 | 0.63 | 0.67 | 5.28 | 2.26 | 0.63 | 0.34 | 0.67 | NA
Results from laboratories 1.06 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 254c (Indicative Value 0.095 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.22 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [15 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.25g | [15 empty] | 3
| 1 | 1 | 0.084 | 0.09525 | < 0.1 | 0.09 | [15 empty] | 1
| 1 | 2 | 0.088 | 0.094334 | < 0.1 | 0.1 | [15 empty] | 4
| 1 | 3 | 0.083 | 0.091713 | < 0.1 | 0.09 | [15 empty] | 16
| 1 | 4 | 0.09 | 0.09581 | < 0.1 | 0.1 | [15 empty] | 0.09507
| 1 | 5 | 0.083 | 0.097798 | < 0.1 | 0.09 | [15 empty] | 10
| 1 | 6 | 0.089 | 0.09593 | < 0.1 | 0.1
| Mean |  | 0.086167 | 0.095139 |  | 0.095
| Median |  | 0.086 | 0.09553 |  | 0.095
| Std Dev. |  | 0.0031885 | 0.0020268 |  | 0.0054772
| Rel.Std.Dev. |  | 3.7% | 2.13% |  | 5.766%
| PDM3 |  | -9.365% | 0.07317% |  | -0.07317%
| Z-Score (Absolute) |  | 0.66 | 0.68 | 6.1 | 0.66
| Analytical results for Mg in OREAS 254c (Certified Value 1.03 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 1.03 | 1.2108 | 1.0216 | 0.93436 | 1.3 | 1.08 | 1.26 | 0.87 | 1.02 | 1.02 | 1.1 | 0.96 | 1.0709 | 1.1 | 1.08 | 1.02 | 0.99 | 1 |  | 1
| 1 | 2 | 1.06 | 1.2381 | 1.0341 | 0.95545 | 1.24 | 1.08 | 1.24 | 0.87 | 1.01 | 0.99 | 1.13 | 0.96 | 1.0672 | 1.09 | 1.04 | 1.04 | 0.98 | 1.004 |  | #N/A
| 1 | 3 | 1.05 | 1.2167 | 1.0233 | 0.92885 | 1.3 | 1.1 | 1.27 | 0.87 | 1.06 | 1 | 1.05 | 0.96 | 1.0294 | 1.08 | 1.08 | 0.97 | 0.97 | 1.006 |  | 16
| 1 | 4 | 1.03 | 1.1951 | 1.0453 | 0.93939 | 1.26 | 1.07 | 1.23 | 0.86 | 0.95 | 1.01 | 1 | 1.02 | 1.0317 | 1.1 | 1.04 | 1.02 | 0.97 | 0.968 |  | 1.0266
| 1 | 5 | 1.04 | 1.2378 | 1.0339 | 0.93775 | 1.25 | 1.08 | 1.24 | 0.86 | 1.05 | 0.99 | 1.11 | 0.99 | 1.0815 | 1.09 | 1.07 | 1 | 0.97 | 0.939 |  | 95
| 1 | 6 | 1.01 | 1.1787 | 1.0434 | 0.94817 | 1.32 | 1.08 | 1.2 | 0.85 | 0.97 | 1.01 | 1.04 | 1.02 | 1.0644 | 1.1 | 1.06 | 1.06 | 0.97 | 0.97
| Mean |  | 1.0367 | 1.2129 | 1.0336 | 0.94066 | 1.2783 | 1.0817 | 1.24 | 0.86333 | 1.01 | 1.0033 | 1.0717 | 0.985 | 1.0575 | 1.0933 | 1.0617 | 1.0183 | 0.975 | 0.98117
| Median |  | 1.035 | 1.2138 | 1.034 | 0.93857 | 1.28 | 1.08 | 1.24 | 0.865 | 1.015 | 1.005 | 1.075 | 0.975 | 1.0658 | 1.095 | 1.065 | 1.02 | 0.97 | 0.985
| Std Dev. |  | 0.017512 | 0.02349 | 0.0098256 | 0.009639 | 0.032506 | 0.0098319 | 0.024495 | 0.008165 | 0.043359 | 0.012111 | 0.049565 | 0.029496 | 0.021685 | 0.008165 | 0.018348 | 0.031252 | 0.0083666 | 0.026716
| Rel.Std.Dev. |  | 1.689% | 1.937% | 0.9506% | 1.025% | 2.543% | 0.909% | 1.975% | 0.9457% | 4.293% | 1.207% | 4.625% | 2.994% | 2.051% | 0.7468% | 1.728% | 3.069% | 0.8581% | 2.723%
| PDM3 |  | 0.9794% | 18.14% | 0.6796% | -8.372% | 24.52% | 5.363% | 20.79% | -15.9% | -1.618% | -2.267% | 4.389% | -4.053% | 3.008% | 6.499% | 3.415% | -0.8064% | -5.027% | -4.427%
| Z-Score (Absolute) |  | 0.02 | 2.48 | 0.02 | 1.32 | 3.39 | 0.65 | 2.86 | 2.4 | 0.35 | 0.44 | 0.51 | 0.7 | 0.31 | 0.81 | 0.37 | 0.23 | 0.84 | 0.75
| Analytical results for Mn in OREAS 254c (Certified Value 0.033 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0326 | 0.034005 | 0.033116 | 0.028836 | 0.03 | 0.0344 | 0.0356 | 0.0288 | 0.0349 | 0.0332 | 0.0328 | 0.0304 | 0.03441 | 0.0347 | 0.0345 | 0.0322 | 0.0314 | 0.0304 |  | 1
| 1 | 2 | 0.0336 | 0.035046 | 0.033374 | 0.029226 | 0.02 | 0.0343 | 0.0352 | 0.0294 | 0.0346 | 0.0319 | 0.0337 | 0.0307 | 0.033531 | 0.0347 | 0.0332 | 0.0327 | 0.031 | 0.0307 |  | 25
| 1 | 3 | 0.0335 | 0.033551 | 0.032968 | 0.028432 | 0.03 | 0.0348 | 0.0365 | 0.0301 | 0.0355 | 0.0328 | 0.0321 | 0.0302 | 0.032915 | 0.0358 | 0.0345 | 0.0306 | 0.0306 | 0.0304 |  | 16
| 1 | 4 | 0.0328 | 0.033174 | 0.033767 | 0.028886 | 0.02 | 0.0343 | 0.035 | 0.0302 | 0.0345 | 0.0329 | 0.0308 | 0.032 | 0.033037 | 0.0354 | 0.033 | 0.0324 | 0.0307 | 0.0294 |  | 0.032684
| 1 | 5 | 0.033 | 0.034677 | 0.033017 | 0.029044 | 0.02 | 0.0344 | 0.035 | 0.0302 | 0.0354 | 0.0321 | 0.0338 | 0.0316 | 0.034477 | 0.0361 | 0.0341 | 0.0318 | 0.031 | 0.0286 |  | 96
| 1 | 6 | 0.0322 | 0.03349 | 0.033477 | 0.029237 | 0.02 | 0.0344 | 0.0342 | 0.0291 | 0.0341 | 0.0331 | 0.0321 | 0.0323 | 0.034423 | 0.0355 | 0.0339 | 0.0333 | 0.0305 | 0.0295
| Mean |  | 0.03295 | 0.03399 | 0.033286 | 0.028943 | 0.023333 | 0.034433 | 0.03525 | 0.029633 | 0.034833 | 0.032667 | 0.03255 | 0.0312 | 0.033799 | 0.035367 | 0.033867 | 0.032167 | 0.030867 | 0.029833
| Median |  | 0.0329 | 0.033778 | 0.033245 | 0.028965 | 0.02 | 0.0344 | 0.0351 | 0.02975 | 0.03475 | 0.03285 | 0.03245 | 0.03115 | 0.03397 | 0.03545 | 0.034 | 0.0323 | 0.03085 | 0.02995
| Std Dev. |  | 0.00053572 | 0.00073415 | 0.00030923 | 0.00030079 | 0.005164 | 0.00018619 | 0.00076354 | 0.00061536 | 0.00054283 | 0.00053914 | 0.0011327 | 0.00088318 | 0.00072866 | 0.00057155 | 0.00064083 | 0.00091797 | 0.00033267 | 0.00080166
| Rel.Std.Dev. |  | 1.626% | 2.16% | 0.929% | 1.039% | 22.13% | 0.5407% | 2.166% | 2.077% | 1.558% | 1.65% | 3.48% | 2.831% | 2.156% | 1.616% | 1.892% | 2.854% | 1.078% | 2.687%
| PDM3 |  | 0.8125% | 3.996% | 1.842% | -11.45% | -28.61% | 5.351% | 7.849% | -9.335% | 6.575% | -0.0544% | -0.4114% | -4.542% | 3.409% | 8.206% | 3.617% | -1.584% | -5.562% | -8.723%
| Z-Score (Absolute) |  | 0.06 | 0.49 | 0.2 | 1.61 | 3.95 | 0.68 | 1.02 | 1.32 | 0.84 | 0.06 | 0.11 | 0.67 | 0.41 | 1.07 | 0.44 | 0.27 | 0.81 | 1.24
| Analytical results for Mo in OREAS 254c (Certified Value 3.06 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 3.3 | 3.1293 | < 10 | 3.3 | 2.96 | 2.54 | 2.9367 | 2.79 | 2.7 | 3.487 | 2.83 | 2.8232 | 3.07 | 2.97 | 3.36 | 2.8 | 3.92 |  |  | 1
| 1 | 2 | 3.4 | 3.0969 | < 10 | 3.3 | 3.05 | 2.51 | 2.9367 | 2.73 | 2.4 | 2.983 | 2.86 | 2.6501 | 3.36 | 2.9 | 3.37 | 2.77 | 3.68 |  |  | 26
| 1 | 3 | 3.3 | 3.156 | < 10 | 3.3 | 3.33 | 2.65 | 3.01 | 2.7 | 3 | 3.685 | 2.87 | 2.5399 | 2.98 | 2.93 | 3.28 | 2.75 | 3.61 |  |  | 16
| 1 | 4 | 3.4 | 3.1496 | < 10 | 3.3 | 3 | 2.6 | 2.83 | 2.74 | 2.9 | 3.569 | 3.09 | 2.497 | 3.02 | 2.8 | 3.43 | 2.79 | 3.75 |  |  | 3.0641
| 1 | 5 | 3.3 | 3.0853 | < 10 | 3.3 | 2.98 | 2.61 | 3 | 2.72 | 3.2 | 3.38 | 3.03 | 2.9176 | 3.35 | 2.94 | 3.39 | 2.78 | 3.78 |  |  | 97
| 1 | 6 | 3.3 | 3.1769 | < 10 | 3.3 | 2.94 | 2.48 | 2.8633 | 2.7 | 2.9 | 3.65 | 3.04 | 2.8453 | 3.19 | 2.91 | 3.45 | 2.76 | 3.66
| Mean |  | 3.3333 | 3.1323 |  | 3.3 | 3.0433 | 2.565 | 2.9294 | 2.73 | 2.85 | 3.459 | 2.9533 | 2.7122 | 3.1617 | 2.9083 | 3.38 | 2.775 | 3.7333
| Median |  | 3.3 | 3.1394 |  | 3.3 | 2.99 | 2.57 | 2.9367 | 2.725 | 2.9 | 3.528 | 2.95 | 2.7367 | 3.13 | 2.92 | 3.38 | 2.775 | 3.715
| Std Dev. |  | 0.05164 | 0.035558 |  | 4.8648e-16 | 0.14542 | 0.065345 | 0.071877 | 0.033466 | 0.27386 | 0.25817 | 0.11219 | 0.17441 | 0.16558 | 0.058452 | 0.06 | 0.018708 | 0.11021
| Rel.Std.Dev. |  | 1.549% | 1.135% |  | 1.474e-14% | 4.778% | 2.548% | 2.454% | 1.226% | 9.609% | 7.464% | 3.799% | 6.43% | 5.237% | 2.01% | 1.775% | 0.6742% | 2.952%
| PDM3 |  | 8.786% | 2.227% |  | 7.698% | -0.6782% | -16.29% | -4.395% | -10.9% | -6.988% | 12.89% | -3.615% | -11.49% | 3.184% | -5.084% | 10.31% | -9.435% | 21.84%
| Z-Score (Absolute) |  | 0.73 | 0.22 | 4.92 | 0.64 | 0 | 1.2 | 0.29 | 0.79 | 0.49 | 1.04 | 0.23 | 0.83 | 0.3 | 0.34 | 0.85 | 0.67 | 1.73
| Analytical results for Na in OREAS 254c (Certified Value 0.335 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.35 | 0.3447 | 0.34538 | 1.28 | 0.34 | 0.356 | 0.27 | 0.33 | 0.34 | 0.336 | 0.32 | 0.31806 | 0.38 | 0.34 | 0.35 | 0.33 | 0.259 |  |  | 1
| 1 | 2 | 0.36 | 0.34942 | 0.35101 | 1.18 | 0.34 | 0.362 | 0.28 | 0.33 | 0.33 | 0.344 | 0.32 | 0.30652 | 0.37 | 0.32 | 0.35 | 0.32 | 0.26 |  |  | #N/A
| 1 | 3 | 0.36 | 0.34293 | 0.33067 | 1.19 | 0.34 | 0.363 | 0.29 | 0.34 | 0.34 | 0.301 | 0.32 | 0.29271 | 0.37 | 0.33 | 0.33 | 0.32 | 0.258 |  |  | 16
| 1 | 4 | 0.36 | 0.35217 | 0.34543 | 1.17 | 0.34 | 0.371 | 0.3 | 0.31 | 0.34 | 0.284 | 0.33 | 0.2979 | 0.38 | 0.32 | 0.35 | 0.32 | 0.248 |  |  | 0.33497
| 1 | 5 | 0.36 | 0.34706 | 0.3491 | 1.18 | 0.34 | 0.366 | 0.29 | 0.34 | 0.33 | 0.314 | 0.33 | 0.32181 | 0.38 | 0.33 | 0.34 | 0.32 | 0.249 |  |  | 98
| 1 | 6 | 0.35 | 0.35323 | 0.34986 | 1.47 | 0.34 | 0.359 | 0.28 | 0.31 | 0.34 | 0.301 | 0.34 | 0.30457 | 0.38 | 0.33 | 0.36 | 0.32 | 0.25
| Mean |  | 0.35667 | 0.34825 | 0.34524 | 1.245 | 0.34 | 0.36283 | 0.285 | 0.32667 | 0.33667 | 0.31333 | 0.32667 | 0.30693 | 0.37667 | 0.32833 | 0.34667 | 0.32167 | 0.254
| Median |  | 0.36 | 0.34824 | 0.34727 | 1.185 | 0.34 | 0.3625 | 0.285 | 0.33 | 0.34 | 0.3075 | 0.325 | 0.30554 | 0.38 | 0.33 | 0.35 | 0.32 | 0.254
| Std Dev. |  | 0.005164 | 0.0040957 | 0.0075096 | 0.11743 | 6.0809e-17 | 0.0052694 | 0.010488 | 0.013663 | 0.005164 | 0.022888 | 0.008165 | 0.01127 | 0.005164 | 0.0075277 | 0.010328 | 0.0040825 | 0.0055498
| Rel.Std.Dev. |  | 1.448% | 1.176% | 2.175% | 9.432% | 1.789e-14% | 1.452% | 3.68% | 4.182% | 1.534% | 7.305% | 2.499% | 3.672% | 1.371% | 2.293% | 2.979% | 1.269% | 2.185%
| PDM3 |  | 6.478% | 3.966% | 3.068% | 271.7% | 1.503% | 8.319% | -14.92% | -2.478% | 0.5074% | -6.458% | -2.478% | -8.37% | 12.45% | -1.98% | 3.493% | -3.971% | -24.17%
| Z-Score (Absolute) |  | 0.9 | 0.52 | 0.39 | 40.83 | 0.15 | 1.18 | 2.32 | 0.45 | 0 | 1.05 | 0.45 | 1.34 | 1.8 | 0.37 | 0.45 | 0.67 | 3.72
| Analytical results for Nb in OREAS 254c (Indicative Value 0.69 ppm) | [20 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.6 | 0.53769 | 26 | 1 | 0.4 | < 1 | < 5 | 0.61 | 0.46 | 1.3624 | 0.9 | 0.68 | 0.8 | 0.44 | 0.4 |  |  |  |  | 1
| 1 | 2 | 0.58 | 0.5567 | 21 | 1 | 0.4 | < 1 | < 5 | 0.58 | 0.51 | 1.1618 | 0.9 | 0.66 | 0.85 | 0.44 | 0.49 |  |  |  |  | 5
| 1 | 3 | 0.62 | 0.53481 | 28 | 1 | 0.4 | < 1 | < 5 | 0.59 | 0.48 | 1.3699 | 0.9 | 0.63 | 0.8 | 0.48 | 0.39 |  |  |  |  | 16
| 1 | 4 | 0.73 | 0.56084 | 23 | 1 | 0.4 | < 1 | < 5 | 0.62 | 0.49 | 1.2668 | 0.9 | 0.62 | 1.35 | 0.45 | 0.4 |  |  |  |  | 0.68502
| 1 | 5 | 0.64 | 0.57169 | 22 | 1 | 0.4 | < 1 | < 5 | 0.6 | 0.46 | 1.306 | 0.9 | 0.71 | 0.94 | 0.42 | 0.41 |  |  |  |  | 11
| 1 | 6 | 0.68 | 0.54527 | 27 | 1.1 | 0.4 | < 1 | < 5 | 0.59 | 0.46 | 1.1976 | 0.9 | 0.67 | 0.88 | 0.45 | 0.38
| Mean |  | 0.64167 | 0.55117 | 24.5 | 1.0167 | 0.4 |  |  | 0.59833 | 0.47667 | 1.2774 | 0.9 | 0.66167 | 0.93667 | 0.44667 | 0.41167
| Median |  | 0.63 | 0.55099 | 24.5 | 1 | 0.4 |  |  | 0.595 | 0.47 | 1.2864 | 0.9 | 0.665 | 0.865 | 0.445 | 0.4
| Std Dev. |  | 0.055287 | 0.014346 | 2.881 | 0.040825 | 6.0809e-17 |  |  | 0.01472 | 0.020656 | 0.085357 | 0 | 0.033116 | 0.20925 | 0.019664 | 0.039707
| Rel.Std.Dev. |  | 8.616% | 2.603% | 11.76% | 4.016% | 1.52e-14% |  |  | 2.46% | 4.333% | 6.682% | 0% | 5.005% | 22.34% | 4.402% | 9.645%
| PDM3 |  | -6.329% | -19.54% | 3477% | 48.41% | -41.61% |  |  | -12.65% | -30.42% | 86.48% | 31.38% | -3.409% | 36.74% | -34.8% | -39.9%
| Z-Score (Absolute) |  | 0 | 0.27 | 69.95 | 1.1 | 0.71 | 0.42 | 5.45 | 0.13 | 0.48 | 1.86 | 0.76 | 0.06 | 0.86 | 0.57 | 0.67
| Analytical results for Nd in OREAS 254c (Indicative Value 21.5 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.09 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 25g | 0.5g | [17 empty] | 1
| 1 | 1 | 21.336 | 21.2 | [17 empty] | 1
| 1 | 2 | 21.719 | 20.6 | [17 empty] | 6
| 1 | 3 | 21.632 | 22.1 | [17 empty] | 16
| 1 | 4 | 22.335 | 21.4 | [17 empty] | 21.5397
| 1 | 5 | 20.998 | 22.6 | [17 empty] | 12
| 1 | 6 | 21.656 | 20.9
| Mean |  | 21.6127 | 21.4667
| Median |  | 21.644 | 21.3
| Std Dev. |  | 0.4449 | 0.75277
| Rel.Std.Dev. |  | 2.059% | 3.507%
| PDM3 |  | 0.3389% | -0.3389%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 254c (Certified Value 76 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 74.6 | 79.9848 | 73.7971 | 67.688 | 64 | 71 | 83.5 | 72.1 | 72.142 | 76 | 74.5 | 77.7 | 78.2382 | 78 | 80 | 78.7 | 75.8 |  |  | 1
| 1 | 2 | 77 | 80.4708 | 74.2833 | 70.156 | 62 | 71 | 79.8 | 74.1 | 72.725 | 74.2 | 76.1 | 77.3 | 77.7448 | 78.7 | 77.2 | 78 | 73.7 |  |  | 29
| 1 | 3 | 76.1 | 76.6908 | 73.2178 | 67.237 | 65 | 73 | 79.2 | 76.9 | 73.5707 | 74.2 | 70.3 | 76.5 | 75.6654 | 79.3 | 79 | 76.5 | 75.3 |  |  | 16
| 1 | 4 | 74.5 | 74.3148 | 75.0333 | 69.782 | 63 | 71 | 81.4 | 76.7 | 72.325 | 75.1 | 70.8 | 83.6 | 76.1826 | 79 | 75 | 81.7 | 75.3 |  |  | 75.7506
| 1 | 5 | 75.8 | 79.2504 | 73.4316 | 73.077 | 63 | 71 | 81.1 | 76.5 | 73.6437 | 74 | 75.6 | 81 | 78.2851 | 80.4 | 78.9 | 80.7 | 74.4 |  |  | 99
| 1 | 6 | 73.6 | 75.9996 | 73.7396 | 71.245 | 62 | 72 | 79.4 | 73.9 | 71.0345 | 74.6 | 68.4 | 82.3 | 75.1011 | 79.6 | 78 | 81 | 74.4
| Mean |  | 75.2667 | 77.7852 | 73.9171 | 69.8642 | 63.1667 | 71.5 | 80.7333 | 75.0333 | 72.5735 | 74.6833 | 72.6167 | 79.7333 | 76.8696 | 79.1667 | 78.0167 | 79.4333 | 74.8167
| Median |  | 75.2 | 77.9706 | 73.7684 | 69.969 | 63 | 71 | 80.45 | 75.3 | 72.525 | 74.4 | 72.65 | 79.35 | 76.9637 | 79.15 | 78.45 | 79.7 | 74.85
| Std Dev. |  | 1.2485 | 2.475 | 0.65584 | 2.189 | 1.169 | 0.83666 | 1.6269 | 1.9582 | 0.97765 | 0.75476 | 3.1946 | 2.9548 | 1.3923 | 0.8165 | 1.7577 | 2.0196 | 0.77825
| Rel.Std.Dev. |  | 1.659% | 3.182% | 0.8873% | 3.133% | 1.851% | 1.17% | 2.015% | 2.61% | 1.347% | 1.011% | 4.399% | 3.706% | 1.811% | 1.031% | 2.253% | 2.542% | 1.04%
| PDM3 |  | -0.6388% | 2.686% | -2.42% | -7.771% | -16.61% | -5.611% | 6.578% | -0.9469% | -4.194% | -1.409% | -4.137% | 5.258% | 1.477% | 4.51% | 2.991% | 4.862% | -1.233%
| Z-Score (Absolute) |  | 0.06 | 0.67 | 0.27 | 1.27 | 2.91 | 0.87 | 1.4 | 0 | 0.6 | 0.09 | 0.59 | 1.15 | 0.45 | 1.01 | 0.73 | 1.08 | 0.05
| Analytical results for P in OREAS 254c (Certified Value 0.052 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0491 | 0.053258 | 0.049214 | 0.06 | 0.0562 | 0.055 | 0.0468 | 0.0482 | 0.053 | 0.053 | 0.05 | 0.05822 | 0.06 | 0.054 | 0.054 | 0.051 | 0.047 |  |  | 1
| 1 | 2 | 0.0511 | 0.053823 | 0.051415 | 0.06 | 0.0557 | 0.054 | 0.0467 | 0.0487 | 0.051 | 0.054 | 0.051 | 0.056608 | 0.06 | 0.053 | 0.054 | 0.051 | 0.047 |  |  | #N/A
| 1 | 3 | 0.051 | 0.053307 | 0.048742 | 0.06 | 0.0562 | 0.053 | 0.0468 | 0.0496 | 0.053 | 0.052 | 0.05 | 0.055921 | 0.06 | 0.054 | 0.051 | 0.051 | 0.047 |  |  | 16
| 1 | 4 | 0.05 | 0.055387 | 0.050847 | 0.06 | 0.0565 | 0.055 | 0.0468 | 0.0483 | 0.053 | 0.049 | 0.052 | 0.056195 | 0.06 | 0.052 | 0.053 | 0.05 | 0.046 |  |  | 0.051768
| 1 | 5 | 0.0501 | 0.05348 | 0.050857 | 0.06 | 0.0561 | 0.053 | 0.0469 | 0.049 | 0.052 | 0.054 | 0.052 | 0.058238 | 0.06 | 0.054 | 0.053 | 0.05 | 0.046 |  |  | 100
| 1 | 6 | 0.0498 | 0.054698 | 0.050976 | 0.06 | 0.0551 | 0.053 | 0.0468 | 0.0478 | 0.052 | 0.052 | 0.053 | 0.056671 | 0.06 | 0.054 | 0.055 | 0.05 | 0.046
| Mean |  | 0.050183 | 0.053992 | 0.050342 | 0.06 | 0.055967 | 0.053833 | 0.0468 | 0.0486 | 0.052333 | 0.052333 | 0.051333 | 0.056975 | 0.06 | 0.0535 | 0.053333 | 0.0505 | 0.0465
| Median |  | 0.05005 | 0.053651 | 0.050852 | 0.06 | 0.05615 | 0.0535 | 0.0468 | 0.0485 | 0.0525 | 0.0525 | 0.0515 | 0.05664 | 0.06 | 0.054 | 0.0535 | 0.0505 | 0.0465
| Std Dev. |  | 0.00075741 | 0.00086522 | 0.0010866 | 0 | 0.00049666 | 0.00098319 | 6.3246e-05 | 0.00064187 | 0.0008165 | 0.0018619 | 0.0012111 | 0.0010092 | 0 | 0.00083666 | 0.0013663 | 0.00054772 | 0.00054772
| Rel.Std.Dev. |  | 1.509% | 1.602% | 2.158% | 0% | 0.8874% | 1.826% | 0.1351% | 1.321% | 1.56% | 3.558% | 2.359% | 1.771% | 0% | 1.564% | 2.562% | 1.085% | 1.178%
| PDM3 |  | -3.062% | 4.295% | -2.755% | 15.9% | 8.11% | 3.989% | -9.597% | -6.12% | 1.091% | 1.091% | -0.8404% | 10.06% | 15.9% | 3.345% | 3.023% | -2.45% | -10.18%
| Z-Score (Absolute) |  | 0.73 | 0.56 | 0.67 | 2.6 | 1.23 | 0.51 | 1.87 | 1.26 | 0 | 0 | 0.34 | 1.57 | 2.6 | 0.4 | 0.34 | 0.62 | 1.98
| Analytical results for Pb in OREAS 254c (Certified Value 8.55 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 8.6 | 9.0171 | < 100 | 13 | 8.8 | 8.8 | 8.2 | 8.3846 | 7.9 | 7.1 | 5.5053 | 7.3 | 9.0367 | 9.6 | 9 | 9 | 8.9 |  |  | 1
| 1 | 2 | 8.8 | 8.5271 | < 100 | 11 | 8.8 | 8.9 | 8.2 | 8.6301 | 7.7 | 7 | 5.7203 | 7.2 | 9.1502 | 9.5 | 8.6 | 9.5 | 8.2 |  |  | 30
| 1 | 3 | 8.6 | 8.558 | < 100 | 13 | 8.6 | 8.1 | 8 | 8.699 | 7.8 | 8.2 | 5.4398 | 7.3 | 9.4463 | 9.4 | 9 | 9.2 | 8.3 |  |  | 16
| 1 | 4 | 8.7 | 9.361 | < 100 | 12 | 8.3 | 7.5 | 8 | 8.5012 | 7.8 | 8.2 | 5.7315 | 8 | 9.8887 | 9.3 | 8.7 | 9.5 | 8.5 |  |  | 8.5537
| 1 | 5 | 8.4 | 8.652 | < 100 | 14 | 8.4 | 8.9 | 8.1 | 8.7118 | 7.8 | 8.8 | 5.3176 | 7.5 | 8.9023 | 9.5 | 9 | 9.5 | 8.5 |  |  | 101
| 1 | 6 | 8.6 | 8.8133 | < 100 | 11 | 8.4 | 7.9 | 7.9 | 8.425 | 7.8 | 7.9 | 5.5241 | 7.7 | 9.7692 | 9.5 | 8.9 | 9.5 | 8.5
| Mean |  | 8.6167 | 8.8214 |  | 12.3333 | 8.55 | 8.35 | 8.0667 | 8.5586 | 7.8 | 7.8667 | 5.5398 | 7.5 | 9.3656 | 9.4667 | 8.8667 | 9.3667 | 8.4833
| Median |  | 8.6 | 8.7327 |  | 12.5 | 8.5 | 8.45 | 8.05 | 8.5657 | 7.8 | 8.05 | 5.5147 | 7.4 | 9.2983 | 9.5 | 8.95 | 9.5 | 8.5
| Std Dev. |  | 0.13292 | 0.32069 |  | 1.2111 | 0.21679 | 0.59917 | 0.12111 | 0.14121 | 0.063246 | 0.69761 | 0.16131 | 0.30332 | 0.40305 | 0.10328 | 0.17512 | 0.21602 | 0.24014
| Rel.Std.Dev. |  | 1.543% | 3.635% |  | 9.819% | 2.536% | 7.176% | 1.501% | 1.65% | 0.8108% | 8.868% | 2.912% | 4.044% | 4.304% | 1.091% | 1.975% | 2.306% | 2.831%
| PDM3 |  | 0.7357% | 3.129% |  | 44.19% | -0.04365% | -2.382% | -5.694% | 0.05696% | -8.812% | -8.032% | -35.24% | -12.32% | 9.491% | 10.67% | 3.658% | 9.504% | -0.823%
| Z-Score (Absolute) |  | 0.06 | 0.26 | 40.39 | 3.68 | 0.01 | 0.2 | 0.48 | 0 | 0.74 | 0.67 | 2.94 | 1.03 | 0.79 | 0.88 | 0.3 | 0.79 | 0.07
| Analytical results for Pd in OREAS 254c (Indicative Value < 10 ppb) | [20 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.16 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 25g | 01g | [17 empty] | 1
| 1 | 1 | < 10 | < 10 | [17 empty] | 1
| 1 | 2 | < 10 | < 10 | [17 empty] | 7
| 1 | 3 | < 10 | < 10 | [17 empty] | 16
| 1 | 4 | < 10 | < 10 | [17 empty] | < 10
| 1 | 5 | < 10 | < 10 | [17 empty] | 13
| 1 | 6 | < 10 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Pr in OREAS 254c (Indicative Value 5.48 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 5.436 | 5.1132 | 5.9 | [16 empty] | 1
| 1 | 2 | 5.501 | 5.1838 | 5.7 | [16 empty] | 8
| 1 | 3 | 5.461 | 5.1242 | 6 | [16 empty] | 16
| 1 | 4 | 5.642 | 5.3459 | 5.6 | [16 empty] | 5.4797
| 1 | 5 | 5.338 | 5.0946 | 5.8 | [16 empty] | 14
| 1 | 6 | 5.438 | 5.257 | 5.7
| Mean |  | 5.4693 | 5.1865 | 5.7833
| Median |  | 5.4495 | 5.154 | 5.75
| Std Dev. |  | 0.10023 | 0.098061 | 0.1472
| Rel.Std.Dev. |  | 1.833% | 1.891% | 2.545%
| PDM3 |  | -0.1893% | -5.352% | 5.541%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.75
| Analytical results for Pt in OREAS 254c (Indicative Value < 5 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.16 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 25g | 01g | [17 empty] | 2
| 1 | 1 | < 5 | < 5 | [17 empty] | 1
| 1 | 2 | < 5 | < 5 | [17 empty] | 9
| 1 | 3 | < 5 | < 5 | [17 empty] | 16
| 1 | 4 | < 5 | < 5 | [17 empty] | < 5
| 1 | 5 | < 5 | < 5 | [17 empty] | 15
| 1 | 6 | < 5 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Rb in OREAS 254c (Certified Value 15.3 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 15.68 | 14.6899 | 14.8 | 15.9 | 14 | 14 | 15.3 | 16.0811 | 19.1 | 16.4 | 16.2 | 14.9 | 15.44 | [6 empty] | 1
| 1 | 2 | 15.91 | 15.1237 | 14.8 | 15.8 | 14.2 | 13.8 | 15 | 15.0947 | 19 | 15.7 | 16.7 | 15.1 | 15.03 | [6 empty] | 32
| 1 | 3 | 16.15 | 14.8093 | 14.8 | 16 | 14.6 | 13.6 | 15.4 | 15.5078 | 19.7 | 16.1 | 15.7 | 14.9 | 15.13 | [6 empty] | 16
| 1 | 4 | 16.6 | 15.3531 | 14.8 | 15.5 | 14.7 | 14 | 16.5 | 15.0111 | 19.4 | 15.2 | 17.3 | 14.7 | 14.84 | [6 empty] | 15.2838
| 1 | 5 | 15.69 | 15.2096 | 14.4 | 15.8 | 14.5 | 13.6 | 16.3 | 16.2103 | 19.7 | 16.4 | 16.9 | 14.7 | 14.54 | [6 empty] | 102
| 1 | 6 | 15.99 | 15.2745 | 14.3 | 15.4 | 13.9 | 13.7 | 16.3 | 15.6275 | 19.3 | 15.9 | 17.2 | 14.7 | 14.62
| Mean |  | 16.0033 | 15.0767 | 14.65 | 15.7333 | 14.3167 | 13.7833 | 15.8 | 15.5887 | 19.3667 | 15.95 | 16.6667 | 14.8333 | 14.9333
| Median |  | 15.95 | 15.1667 | 14.8 | 15.8 | 14.35 | 13.75 | 15.85 | 15.5676 | 19.35 | 16 | 16.8 | 14.8 | 14.935
| Std Dev. |  | 0.34326 | 0.26702 | 0.23452 | 0.23381 | 0.33116 | 0.18348 | 0.63875 | 0.49285 | 0.29439 | 0.45935 | 0.61536 | 0.1633 | 0.33655
| Rel.Std.Dev. |  | 2.145% | 1.771% | 1.601% | 1.486% | 2.313% | 1.331% | 4.043% | 3.162% | 1.52% | 2.88% | 3.692% | 1.101% | 2.254%
| PDM3 |  | 4.708% | -1.355% | -4.147% | 2.941% | -6.328% | -9.817% | 3.378% | 1.995% | 26.71% | 4.359% | 9.048% | -2.947% | -2.293%
| Z-Score (Absolute) |  | 0.43 | 0.53 | 0.97 | 0.15 | 1.31 | 1.86 | 0.22 | 0 | 3.89 | 0.37 | 1.11 | 0.78 | 0.67
| Analytical results for Re in OREAS 254c (Certified Value < 0.001 ppm) | [20 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 1
| 1 | 2 | < 0.001 | < 0.05 | 0.002 | < 0.002 | < 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 33
| 1 | 3 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 16
| 1 | 4 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | < 0.001
| 1 | 5 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 103
| 1 | 6 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001
| Mean |  |  |  | 0.002 |  |  | 0.001
| Median |  |  |  | 0.002 |  |  | 0.001
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 395.82 | 3.37 | 7.42 | 0.67 | 0.67 | 395.82 | 0.67 | 0.67 | 0.67
| Analytical results for S in OREAS 254c (Certified Value 0.02 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0186 | 0.018733 |  | 0.56 | 0.02 | < 0.001 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.018016 | 0.03 | 0.019765 | 0.02 | 0.02 | 0.03 | 0.03 | 0.02 |  | 1
| 1 | 2 | 0.0191 | 0.018448 |  | 0.55 | 0.02 | < 0.001 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.018626 | 0.03 | 0.020168 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |  | 18
| 1 | 3 | 0.0186 | 0.018809 |  | 0.54 | 0.02 | < 0.001 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.018513 | 0.03 | 0.020441 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 |  | 16
| 1 | 4 | 0.0182 | 0.019539 |  | 0.54 | 0.02 | < 0.001 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.017822 | 0.02 | 0.019899 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 |  | 0.019675
| 1 | 5 | 0.0188 | 0.01851 |  | 0.54 | 0.02 | < 0.001 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.017916 | 0.03 | 0.020096 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |  | 104
| 1 | 6 | 0.0184 | 0.0194 |  | 0.68 | 0.02 | < 0.001 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.018427 | 0.03 | 0.019842 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02
| Mean |  | 0.018617 | 0.018907 |  | 0.56833 | 0.02 |  | 0.02 | 0.02 | 0.02 |  | 0.01822 | 0.028333 | 0.020035 | 0.02 | 0.018333 | 0.023333 | 0.023333 | 0.02
| Median |  | 0.0186 | 0.018771 |  | 0.545 | 0.02 |  | 0.02 | 0.02 | 0.02 |  | 0.018221 | 0.03 | 0.019997 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02
| Std Dev. |  | 0.00031252 | 0.00045852 |  | 0.055287 | 0 |  | 0 | 0 | 0 |  | 0.00034227 | 0.0040825 | 0.00025101 | 0 | 0.0040825 | 0.005164 | 0.005164 | 0
| Rel.Std.Dev. |  | 1.679% | 2.425% |  | 9.728% | 0% |  | 0% | 0% | 0% |  | 1.879% | 14.41% | 1.253% | 0% | 22.27% | 22.13% | 22.13% | 0%
| PDM3 |  | -5.38% | -3.907% |  | 2789% | 1.651% |  | 1.651% | 1.651% | 1.651% |  | -7.397% | 44% | 1.829% | 1.651% | -6.82% | 18.59% | 18.59% | 1.651%
| Z-Score (Absolute) |  | 0.67 | 0.53 | NA | 267.29 | 0 | 9.51 | 0 | 0 | 0 | 2.44 | 0.87 | 4.06 | 0.02 | 0 | 0.81 | 1.62 | 1.62 | 0
| Analytical results for Sb in OREAS 254c (Certified Value 0.56 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 0.71 | 0.54077 | < 10 | < 1 | 0.61 | 0.56 | 0.33 | < 5 | 0.54 | 0.6 | 0.492 | 0.45 | 0.63262 | 0.72 | 0.54 | 0.62 | 0.53 |  |  | 1
| 1 | 2 | 0.71 | 0.53447 | < 10 | < 1 | 0.55 | 0.56 | 0.34 | < 5 | 0.55 | 0.6 | 0.468 | 0.45 | 0.63082 | 0.84 | 0.55 | 0.65 | 0.53 |  |  | 34
| 1 | 3 | 0.72 | 0.54576 | < 10 | < 1 | 0.56 | 0.57 | 0.36 | < 5 | 0.52 | 0.6 | 0.474 | 0.44 | 0.6516 | 0.72 | 0.53 | 0.58 | 0.52 |  |  | 16
| 1 | 4 | 0.69 | 0.53792 | < 10 | < 1 | 0.58 | 0.57 | 0.36 | < 5 | 0.55 | 0.6 | 0.505 | 0.45 | 0.62575 | 0.72 | 0.51 | 0.66 | 0.53 |  |  | 0.5577
| 1 | 5 | 0.72 | 0.56308 | < 10 | 3.8 | 0.55 | 0.55 | 0.35 | < 5 | 0.55 | 0.6 | 0.489 | 0.44 | 0.68329 | 0.77 | 0.55 | 0.6 | 0.51 |  |  | 105
| 1 | 6 | 0.73 | 0.55333 | < 10 | 1.2 | 0.5 | 0.55 | 0.33 | < 5 | 0.53 | 0.6 | 0.428 | 0.48 | 0.63713 | 0.75 | 0.53 | 0.83 | 0.52
| Mean |  | 0.71333 | 0.54589 |  | 2.5 | 0.55833 | 0.56 | 0.345 |  | 0.54 | 0.6 | 0.476 | 0.45167 | 0.64354 | 0.75333 | 0.535 | 0.65667 | 0.52333
| Median |  | 0.715 | 0.54327 |  | 2.5 | 0.555 | 0.56 | 0.345 |  | 0.545 | 0.6 | 0.4815 | 0.45 | 0.63488 | 0.735 | 0.535 | 0.635 | 0.525
| Std Dev. |  | 0.013663 | 0.010689 |  | 1.8385 | 0.03656 | 0.0089443 | 0.013784 |  | 0.012649 | 0 | 0.026974 | 0.01472 | 0.021371 | 0.047188 | 0.015166 | 0.090037 | 0.008165
| Rel.Std.Dev. |  | 1.915% | 1.958% |  | 73.54% | 6.548% | 1.597% | 3.995% |  | 2.342% | 0% | 5.667% | 3.259% | 3.321% | 6.264% | 2.835% | 13.71% | 1.56%
| PDM3 |  | 27.91% | -2.118% |  | 348.3% | 0.1137% | 0.4125% | -38.14% |  | -3.174% | 7.585% | -14.65% | -19.01% | 15.39% | 35.08% | -4.07% | 17.75% | -6.162%
| Z-Score (Absolute) |  | 1.23 | 0.1 | 35.15 | NA | 0 | 0.01 | 1.69 | 15.37 | 0.15 | NA | 0.65 | 0.84 | 0.67 | 1.54 | 0.18 | 0.78 | 0.28
Results from laboratories 1.06 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 254c (Certified Value 3.92 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 3.9 | 3.7173 | 4 | 4.1 | 3.5 | < 5 | 3.7 | 3.7 | 3.4 | 3.7326 | 4.5 | 4.3 | 3.9 | 3.8 | 4.7 |  |  |  |  | 1
| 1 | 2 | 4 | 3.8004 | 4 | 4 | 3.5 | < 5 | 3.6 | 3.8 | 3.4 | 3.8357 | 4.7 | 4.1 | 4 | 3.8 | 4.6 |  |  |  |  | 35
| 1 | 3 | 4 | 3.6779 | 4 | 4.3 | 3.5 | < 5 | 3.7 | 3.5 | 3.5 | 3.3682 | 4.6 | 4.3 | 3.8 | 3.8 | 4.7 |  |  |  |  | 16
| 1 | 4 | 3.9 | 3.8389 | 4 | 3.8 | 3.6 | < 5 | 3.8 | 3.4 | 3.7 | 3.4252 | 4.4 | 4 | 4.1 | 3.9 | 4.6 |  |  |  |  | 3.9183
| 1 | 5 | 4 | 3.755 | 4 | 4.1 | 3.5 | < 5 | 3.7 | 3.5 | 3.6 | 3.8462 | 4.7 | 4.2 | 4.3 | 3.7 | 4.6 |  |  |  |  | 106
| 1 | 6 | 3.9 | 3.8227 | 4 | 3.9 | 3.5 | < 5 | 3.7 | 3.4 | 3.7 | 3.61 | 4.6 | 4.2 | 4.2 | 3.8 | 4.5
| Mean |  | 3.95 | 3.7687 | 4 | 4.0333 | 3.5167 |  | 3.7 | 3.55 | 3.55 | 3.6363 | 4.5833 | 4.1833 | 4.05 | 3.8 | 4.6167
| Median |  | 3.95 | 3.7777 | 4 | 4.05 | 3.5 |  | 3.7 | 3.5 | 3.55 | 3.6713 | 4.6 | 4.2 | 4.05 | 3.8 | 4.6
| Std Dev. |  | 0.054772 | 0.063123 | 0 | 0.17512 | 0.040825 |  | 0.063246 | 0.16432 | 0.13784 | 0.20506 | 0.1169 | 0.1169 | 0.18708 | 0.063246 | 0.075277
| Rel.Std.Dev. |  | 1.387% | 1.675% | 0% | 4.342% | 1.161% |  | 1.709% | 4.629% | 3.883% | 5.639% | 2.551% | 2.795% | 4.619% | 1.664% | 1.631%
| PDM3 |  | 0.8081% | -3.818% | 2.084% | 2.935% | -10.25% |  | -5.572% | -9.4% | -9.4% | -7.198% | 16.97% | 6.763% | 3.36% | -3.02% | 17.82%
| Z-Score (Absolute) |  | 0.46 | 0.04 | NA | 0.69 | 0.75 | 3.58 | 0.24 | 0.65 | 0.65 | 0.41 | 2.23 | 1.11 | 0.74 | 0.04 | 2.32
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 254c (Indicative Value 0.28 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | < 1.5 | < 1 | 0.4 | 0.3 | < 5 | 0.3 | < 0.5 | 0.104 | 0.3 | < 1 | 1.4 | 0.2 | < 0.2 | < 0.2 |  |  |  |  | 1
| 1 | 2 | < 1 | < 1.5 | < 1 | 0.5 | 0.3 | < 5 | 0.5 | < 0.5 | 0.103 | 0.3 | < 1 | 1.3 | < 0.2 | < 0.2 | 0.2 |  |  |  |  | 1
| 1 | 3 | < 1 | < 1.5 | < 1 | 0.3 | 0.4 | < 5 | 0.3 | < 0.5 | 0.102 | 0.2 | < 1 | 1.3 | 0.2 | < 0.2 | 0.2 |  |  |  |  | 16
| 1 | 4 | < 1 | < 1.5 | < 1 | 0.5 | 0.4 | < 5 | 0.4 | < 0.5 | 0.105 | 0.3 | < 1 | 1.4 | < 0.2 | < 0.2 | 0.2 |  |  |  |  | 0.27669
| 1 | 5 | < 1 | < 1.5 | < 1 | 0.4 | 0.4 | < 5 | 0.4 | < 0.5 | 0.104 | 0.4 | < 1 | 1.3 | < 0.2 | < 0.2 | 0.2 |  |  |  |  | 10
| 1 | 6 | < 1 | < 1.5 | < 1 | 0.4 | 0.4 | < 5 | 0.3 | < 0.5 | 0.103 | 0.2 | < 1 | 1.3 | 0.2 | < 0.2 | 0.2
| Mean |  |  |  |  | 0.41667 | 0.36667 |  | 0.36667 |  | 0.1035 | 0.28333 |  | 1.3333 | 0.2 |  | 0.2
| Median |  |  |  |  | 0.4 | 0.4 |  | 0.35 |  | 0.1035 | 0.3 |  | 1.3 | 0.2 |  | 0.2
| Std Dev. |  |  |  |  | 0.075277 | 0.05164 |  | 0.08165 |  | 0.0010488 | 0.075277 |  | 0.05164 | 3.3993e-17 |  | 0
| Rel.Std.Dev. |  |  |  |  | 18.07% | 14.08% |  | 22.27% |  | 1.013% | 26.57% |  | 3.873% | 1.7e-14% |  | 0%
| PDM3 |  |  |  |  | 50.59% | 32.52% |  | 32.52% |  | -62.59% | 2.401% |  | 381.9% | -27.72% |  | -27.72%
| Z-Score (Absolute) |  | 0.67 | 1.94 | 0.67 | 0.25 | 0 | 10.79 | 0 | 0.59 | 1.33 | 0.42 | 0.67 | 4.89 | 1.1 | 1.35 | 0.93
| Analytical results for Sm in OREAS 254c (Indicative Value 3.78 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 3.821 | 3.7334 | 3.9 | [16 empty] | 1
| 1 | 2 | 3.791 | 3.7514 | 4.7 | [16 empty] | 5
| 1 | 3 | 3.778 | 3.7568 | 4.4 | [16 empty] | 16
| 1 | 4 | 3.881 | 3.854 | 4 | [16 empty] | 3.7828
| 1 | 5 | 3.713 | 3.8187 | 3.4 | [16 empty] | 11
| 1 | 6 | 3.745 | 3.7508 | 4.1
| Mean |  | 3.7882 | 3.7775 | 4.0833
| Median |  | 3.7845 | 3.7541 | 4.05
| Std Dev. |  | 0.058857 | 0.047566 | 0.4446
| Rel.Std.Dev. |  | 1.554% | 1.259% | 10.89%
| PDM3 |  | 0.1408% | -0.1408% | 7.944%
| Z-Score (Absolute) |  | 0 | 0.67 | 18.69
| Analytical results for Sn in OREAS 254c (Certified Value 1.26 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 1.28 | < 2.5 | 7 | 1.6 | 1.47 | 1 | < 10 | 1.2 | 1.1 | 1.2875 | 1.5 | 1.2 | 1.3 | 1.1 | 1.2 |  |  |  |  | 1
| 1 | 2 | 1.23 | < 2.5 | < 1 | 1.5 | 1.48 | 1.1 | < 10 | 1.2 | 1.1 | 1.0489 | 1.6 | 1.2 | 1.4 | 1.1 | 1.2 |  |  |  |  | 21
| 1 | 3 | 1.33 | < 2.5 | 3 | 1.6 | 1.55 | 1.1 | < 10 | 1.2 | 1.1 | 1.2551 | 1.6 | 1.2 | 1.3 | 1.1 | 1.2 |  |  |  |  | 16
| 1 | 4 | 1.35 | < 2.5 | 1 | 1.5 | 1.48 | 1.2 | < 10 | 1.2 | 1.2 | 1.0618 | 1.6 | 1.1 | 1.8 | 1.1 | 1.3 |  |  |  |  | 1.2599
| 1 | 5 | 1.25 | < 2.5 | 3 | 1.5 | 1.46 | 1.1 | < 10 | 1.2 | 1.2 | 1.3182 | 1.6 | 1.2 | 1.5 | 1.1 | 1.3 |  |  |  |  | 107
| 1 | 6 | 1.3 | < 2.5 | 5 | 1.6 | 1.44 | 1.1 | < 10 | 1.2 | 1.2 | 1.1473 | 1.6 | 1.2 | 1.5 | 1.1 | 1.2
| Mean |  | 1.29 |  | 3.8 | 1.55 | 1.48 | 1.1 |  | 1.2 | 1.15 | 1.1865 | 1.5833 | 1.1833 | 1.4667 | 1.1 | 1.2333
| Median |  | 1.29 |  | 3 | 1.55 | 1.475 | 1.1 |  | 1.2 | 1.15 | 1.2012 | 1.6 | 1.2 | 1.45 | 1.1 | 1.2
| Std Dev. |  | 0.046043 |  | 2.2804 | 0.054772 | 0.037417 | 0.063246 |  | 0 | 0.054772 | 0.11683 | 0.040825 | 0.040825 | 0.18619 | 2.4324e-16 | 0.05164
| Rel.Std.Dev. |  | 3.569% |  | 60.01% | 3.534% | 2.528% | 5.75% |  | 0% | 4.763% | 9.847% | 2.578% | 3.45% | 12.69% | 2.211e-14% | 4.187%
| PDM3 |  | 2.387% |  | 201.6% | 23.02% | 17.47% | -12.69% |  | -4.756% | -8.724% | -5.83% | 25.67% | -6.079% | 16.41% | -12.69% | -2.11%
| Z-Score (Absolute) |  | 0.28 | 0.05 | NA | 1.78 | 1.38 | 0.82 | 21.72 | 0.24 | 0.53 | 0.32 | 1.97 | 0.34 | 1.3 | 0.82 | 0.05
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 254c (Certified Value 60 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  | 0
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 59.87 | 59.8858 | 226 | 58.6 | 70.8 | 61 | 60.93 | 62 | 57.6 | 60.6 | 55.7709 | 56.4 | 62.1 | 62.4 | 59.5 | 46.3 |  |  |  | 1
| 1 | 2 | 60.24 | 60.3954 | 220 | 56.5 | 72.5 | 61 | 60.7067 | 60.5 | 59.1 | 60 | 53.7302 | 56.7 | 59.3 | 64.9 | 58.3 | 46.7 |  |  |  | 22
| 1 | 3 | 60.27 | 59.1485 | 227 | 56 | 74.9 | 61 | 62.6533 | 60 | 55.9 | 60.3 | 54.5868 | 56.8 | 61.1 | 62.5 | 58 | 45.9 |  |  |  | 16
| 1 | 4 | 61.7 | 60.9627 | 220 | 57.5 | 71.2 | 61 | 58.13 | 60.4 | 53.4 | 65.4 | 54.0295 | 56.9 | 57.8 | 65.9 | 59.1 | 44.9 |  |  |  | 59.5819
| 1 | 5 | 59.21 | 60.1643 | 218 | 57.1 | 70 | 61 | 61.6867 | 59.7 | 59.2 | 63.4 | 55.8072 | 57.9 | 61.2 | 67.6 | 58.6 | 43.9 |  |  |  | 108
| 1 | 6 | 59.79 | 61.1319 | 265 | 57.3 | 71.3 | 60 | 58.4133 | 60.3 | 56.1 | 63.9 | 55.5734 | 57.5 | 60.6 | 67.4 | 58.3 | 44.8
| Mean |  | 60.18 | 60.2814 | 229.3333 | 57.1667 | 71.7833 | 60.8333 | 60.42 | 60.4833 | 56.8833 | 62.2667 | 54.9163 | 57.0333 | 60.35 | 65.1167 | 58.6333 | 45.4167
| Median |  | 60.055 | 60.2799 | 223 | 57.2 | 71.25 | 61 | 60.8183 | 60.35 | 56.85 | 62 | 55.0801 | 56.85 | 60.85 | 65.4 | 58.45 | 45.4
| Std Dev. |  | 0.838 | 0.72861 | 17.84 | 0.89368 | 1.7291 | 0.40825 | 1.8003 | 0.79854 | 2.214 | 2.2607 | 0.9228 | 0.55737 | 1.5502 | 2.2921 | 0.56451 | 1.0591
| Rel.Std.Dev. |  | 1.392% | 1.209% | 7.779% | 1.563% | 2.409% | 0.6711% | 2.98% | 1.32% | 3.892% | 3.631% | 1.68% | 0.9773% | 2.569% | 3.52% | 0.9628% | 2.332%
| PDM3 |  | 1.004% | 1.174% | 284.9% | -4.054% | 20.48% | 2.1% | 1.407% | 1.513% | -4.529% | 4.506% | -7.83% | -4.277% | 1.289% | 9.289% | -1.592% | -23.77%
| Z-Score (Absolute) |  | 0.04 | 0.01 | 44.69 | 0.83 | 3.03 | 0.14 | 0.03 | 0.04 | 0.91 | 0.52 | 1.43 | 0.87 | 0.01 | 1.27 | 0.44 | 3.94
| Analytical results for Ta in OREAS 254c (Certified Value < 0.01 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 23
| 1 | 3 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 109
| 1 | 6 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01
| Mean | [9 empty] | 0.3 |  | 0.01
| Median | [9 empty] | 0.3 |  | 0.01
| Std Dev. | [9 empty] | 0 |  | 0
| Rel.Std.Dev. | [9 empty] | 0% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 7.42 | 7.42 | 0.67 | 1008.77 | 0.67 | 0.67 | 7.42 | 118.68 | 0.67 | 0.67 | 0.67
| Analytical results for Tb in OREAS 254c (Certified Value 0.44 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.17 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.25g | [14 empty] | 3
| 1 | 1 | 0.416 | 0.4207 | 0.4 | 0.5 | 0.45 | [14 empty] | 1
| 1 | 2 | 0.414 | 0.42184 | 0.4 | 0.5 | 0.47 | [14 empty] | 24
| 1 | 3 | 0.416 | 0.42535 | 0.4 | 0.5 | 0.46 | [14 empty] | 16
| 1 | 4 | 0.426 | 0.43437 | 0.4 | 0.5 | 0.47 | [14 empty] | 0.44055
| 1 | 5 | 0.407 | 0.42143 | 0.4 | 0.5 | 0.47 | [14 empty] | 110
| 1 | 6 | 0.409 | 0.44615 | 0.4 | 0.5 | 0.46
| Mean |  | 0.41467 | 0.42831 | 0.4 | 0.5 | 0.46333
| Median |  | 0.415 | 0.42359 | 0.4 | 0.5 | 0.465
| Std Dev. |  | 0.0066833 | 0.01011 | 6.0809e-17 | 0 | 0.008165
| Rel.Std.Dev. |  | 1.612% | 2.36% | 1.52e-14% | 0% | 1.762%
| PDM3 |  | -5.875% | -2.779% | -9.204% | 13.5% | 5.172%
| Z-Score (Absolute) |  | 0.32 | 0 | 0.67 | 1.71 | 0.83
| Analytical results for Te in OREAS 254c (Indicative Value 0.04 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 3
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.1 | < 0.1 | < 0.05 | 0.03 | 0.05 | < 5 | 0.04 | < 0.2 | < 0.03 | 0.03 | < 0.05 | 0.07 | 0.03 | 0.02 | 0.02 |  |  |  |  | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.05 | 0.07 | 0.04 | < 5 | 0.03 | < 0.2 | < 0.03 | 0.03 | < 0.05 | 0.07 | 0.03 | 0.05 | 0.04 |  |  |  |  | 2
| 1 | 3 | < 0.1 | < 0.1 | < 0.05 | 0.07 | 0.04 | < 5 | 0.03 | < 0.2 | < 0.03 | 0.01 | < 0.05 | 0.05 | 0.04 | < 0.01 | 0.05 |  |  |  |  | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.05 | 0.03 | 0.05 | < 5 | 0.04 | < 0.2 | < 0.03 | 0.02 | < 0.05 | 0.06 | 0.03 | 0.02 | 0.03 |  |  |  |  | 0.040185
| 1 | 5 | < 0.1 | < 0.1 | 0.05 | 0.04 | 0.04 | < 5 | 0.04 | < 0.2 | < 0.03 | 0.03 | < 0.05 | 0.08 | 0.03 | < 0.01 | 0.03 |  |  |  |  | 12
| 1 | 6 | < 0.1 | < 0.1 | 0.05 | 0.06 | 0.05 | < 5 | 0.03 | < 0.2 | < 0.03 | 0.03 | < 0.05 | 0.06 | 0.04 | 0.01 | 0.03
| Mean |  |  |  | 0.05 | 0.05 | 0.045 |  | 0.035 |  |  | 0.025 |  | 0.065 | 0.033333 | 0.025 | 0.033333
| Median |  |  |  | 0.05 | 0.05 | 0.045 |  | 0.035 |  |  | 0.03 |  | 0.065 | 0.03 | 0.02 | 0.03
| Std Dev. |  |  |  | 0 | 0.018974 | 0.0054772 |  | 0.0054772 |  |  | 0.0083666 |  | 0.010488 | 0.005164 | 0.017321 | 0.010328
| Rel.Std.Dev. |  |  |  | 0% | 37.95% | 12.17% |  | 15.65% |  |  | 33.47% |  | 16.14% | 15.49% | 69.28% | 30.98%
| PDM3 |  |  |  | 24.42% | 24.42% | 11.98% |  | -12.9% |  |  | -37.79% |  | 61.75% | -17.05% | -37.79% | -17.05%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.07 | 0.67 | 0.45 | 110.81 | 0 | 2.92 | 0.9 | 0.45 | 0.45 | 1.35 | 0.07 | 0.75 | 0.07
| Analytical results for Th in OREAS 254c (Certified Value 5.62 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 5.69 | 5.4796 | < 100 | 5.53 | 5.9 | < 10 | 5.3 | 5.5 | 4.7 | 6.3881 | 5.54 | 5.7 | 6.1 | 5.4 | 6.68 |  |  |  |  | 1
| 1 | 2 | 5.86 | 5.6229 | < 100 | 5.62 | 6.1 | < 10 | 5.2 | 5.5 | 4.7 | 6.4976 | 5.34 | 5.4 | 6.3 | 5.4 | 6.45 |  |  |  |  | 26
| 1 | 3 | 5.8 | 5.4816 | < 100 | 5.47 | 6.1 | < 10 | 5.2 | 5.6 | 4.7 | 6.3367 | 4.93 | 5.8 | 6 | 5.3 | 6.83 |  |  |  |  | 16
| 1 | 4 | 5.89 | 5.6407 | < 100 | 5.3 | 5.9 | < 10 | 5.3 | 5.6 | 5.1 | 6.36 | 5.22 | 5.5 | 6.2 | 5.4 | 6.64 |  |  |  |  | 5.6198
| 1 | 5 | 5.62 | 5.5393 | < 100 | 5.3 | 6.1 | < 10 | 5.2 | 6 | 4.9 | 6.4642 | 5.49 | 5.7 | 6.5 | 5.4 | 6.65 |  |  |  |  | 111
| 1 | 6 | 5.73 | 5.6885 | < 100 | 5.44 | 5.8 | < 10 | 5.2 | 5.3 | 5 | 6.473 | 5.48 | 5.6 | 6.3 | 5.5 | 6.56
| Mean |  | 5.765 | 5.5754 |  | 5.4433 | 5.9833 |  | 5.2333 | 5.5833 | 4.85 | 6.4199 | 5.3333 | 5.6167 | 6.2333 | 5.4 | 6.635
| Median |  | 5.765 | 5.5811 |  | 5.455 | 6 |  | 5.2 | 5.55 | 4.8 | 6.4262 | 5.41 | 5.65 | 6.25 | 5.4 | 6.645
| Std Dev. |  | 0.10368 | 0.087876 |  | 0.12691 | 0.13292 |  | 0.05164 | 0.23166 | 0.17607 | 0.066859 | 0.22993 | 0.1472 | 0.17512 | 0.063246 | 0.12661
| Rel.Std.Dev. |  | 1.798% | 1.576% |  | 2.332% | 2.221% |  | 0.9867% | 4.149% | 3.63% | 1.041% | 4.311% | 2.621% | 2.809% | 1.171% | 1.908%
| PDM3 |  | 2.585% | -0.7884% |  | -3.139% | 6.47% |  | -6.876% | -0.6481% | -13.7% | 14.24% | -5.097% | -0.05493% | 10.92% | -3.91% | 18.07%
| Z-Score (Absolute) |  | 0.35 | 0.02 | 85.57 | 0.27 | 0.77 | 1.12 | 0.67 | 0 | 1.41 | 1.61 | 0.48 | 0.06 | 1.25 | 0.35 | 2.03
| Analytical results for Ti in OREAS 254c (Certified Value 0.147 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.1483 | 0.15079 | 0.12184 | 0.16 | 0.178 | 0.2 | 0.14 | 0.15 | 0.148 | 0.137 | 0.145 | 0.23009 | 0.19 | 0.166 | 0.153 | 0.148 | 0.124 |  |  | 1
| 1 | 2 | 0.1522 | 0.15273 | 0.12149 | 0.14 | 0.18 | 0.19 | 0.14 | 0.15 | 0.142 | 0.142 | 0.141 | 0.23221 | 0.18 | 0.158 | 0.155 | 0.144 | 0.125 |  |  | 27
| 1 | 3 | 0.158 | 0.15088 | 0.11446 | 0.17 | 0.184 | 0.19 | 0.14 | 0.15 | 0.145 | 0.145 | 0.142 | 0.22906 | 0.18 | 0.166 | 0.143 | 0.141 | 0.122 |  |  | 16
| 1 | 4 | 0.1559 | 0.15455 | 0.12272 | 0.16 | 0.18 | 0.19 | 0.14 | 0.14 | 0.146 | 0.141 | 0.147 | 0.22907 | 0.19 | 0.159 | 0.158 | 0.142 | 0.119 |  |  | 0.14718
| 1 | 5 | 0.1496 | 0.15202 | 0.12307 | 0.14 | 0.181 | 0.19 | 0.14 | 0.15 | 0.142 | 0.15 | 0.147 | 0.23077 | 0.19 | 0.165 | 0.15 | 0.141 | 0.116 |  |  | 112
| 1 | 6 | 0.1485 | 0.15396 | 0.12345 | 0.16 | 0.181 | 0.19 | 0.13 | 0.14 | 0.143 | 0.145 | 0.151 | 0.23218 | 0.19 | 0.161 | 0.159 | 0.143 | 0.12
| Mean |  | 0.15208 | 0.15249 | 0.12117 | 0.155 | 0.18067 | 0.19167 | 0.13833 | 0.14667 | 0.14433 | 0.14333 | 0.1455 | 0.23056 | 0.18667 | 0.1625 | 0.153 | 0.14317 | 0.121
| Median |  | 0.1509 | 0.15237 | 0.12228 | 0.16 | 0.1805 | 0.19 | 0.14 | 0.15 | 0.144 | 0.1435 | 0.146 | 0.23043 | 0.19 | 0.163 | 0.154 | 0.1425 | 0.121
| Std Dev. |  | 0.0040721 | 0.0015626 | 0.0033709 | 0.012247 | 0.0019664 | 0.0040825 | 0.0040825 | 0.005164 | 0.0024221 | 0.0044121 | 0.0036742 | 0.0014199 | 0.005164 | 0.0036194 | 0.0058992 | 0.0026394 | 0.0033466
| Rel.Std.Dev. |  | 2.678% | 1.025% | 2.782% | 7.902% | 1.088% | 2.13% | 2.951% | 3.521% | 1.678% | 3.078% | 2.525% | 0.6158% | 2.766% | 2.227% | 3.856% | 1.844% | 2.766%
| PDM3 |  | 3.328% | 3.603% | -17.67% | 5.31% | 22.75% | 30.22% | -6.014% | -0.3519% | -1.937% | -2.617% | -1.145% | 56.65% | 26.82% | 10.41% | 3.951% | -2.73% | -17.79%
| Z-Score (Absolute) |  | 0 | 0.03 | 2.34 | 0.22 | 2.16 | 2.99 | 1.04 | 0.41 | 0.59 | 0.66 | 0.5 | 5.93 | 2.62 | 0.79 | 0.07 | 0.67 | 2.35
| Analytical results for Tl in OREAS 254c (Certified Value 0.11 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.11 | 0.10754 | 0.12 | 0.13 | 0.09 | < 10 | 0.1 | < 0.1 | 0.11 | 0.09 | 0.10443 | 0.15 | 0.11 | 0.12 | 0.1 |  |  |  |  | 1
| 1 | 2 | 0.12 | 0.10584 | 0.11 | 0.13 | 0.11 | < 10 | 0.1 | < 0.1 | 0.107 | 0.09 | 0.091997 | 0.17 | 0.1 | 0.12 | 0.1 |  |  |  |  | 28
| 1 | 3 | 0.11 | 0.11467 | 0.12 | 0.13 | 0.09 | < 10 | 0.1 | < 0.1 | 0.111 | 0.09 | 0.099709 | 0.15 | 0.1 | 0.12 | 0.1 |  |  |  |  | 16
| 1 | 4 | 0.11 | 0.11334 | 0.12 | 0.13 | 0.1 | < 10 | 0.11 | < 0.1 | 0.1094 | 0.1 | 0.10006 | 0.15 | 0.1 | 0.13 | 0.1 |  |  |  |  | 0.10789
| 1 | 5 | 0.11 | 0.12576 | 0.12 | 0.13 | 0.1 | < 10 | 0.1 | < 0.1 | 0.107 | 0.1 | 0.10129 | 0.16 | 0.1 | 0.12 | 0.1 |  |  |  |  | 113
| 1 | 6 | 0.11 | 0.10933 | 0.11 | 0.13 | 0.09 | < 10 | 0.1 | < 0.1 | 0.1077 | 0.1 | 0.091893 | 0.16 | 0.1 | 0.12 | 0.1
| Mean |  | 0.11167 | 0.11275 | 0.11667 | 0.13 | 0.096667 |  | 0.10167 |  | 0.10868 | 0.095 | 0.098231 | 0.15667 | 0.10167 | 0.12167 | 0.1
| Median |  | 0.11 | 0.11133 | 0.12 | 0.13 | 0.095 |  | 0.1 |  | 0.10855 | 0.095 | 0.099886 | 0.155 | 0.1 | 0.12 | 0.1
| Std Dev. |  | 0.0040825 | 0.0072097 | 0.005164 | 0 | 0.008165 |  | 0.0040825 |  | 0.0016881 | 0.0054772 | 0.0051451 | 0.008165 | 0.0040825 | 0.0040825 | 1.5202e-17
| Rel.Std.Dev. |  | 3.656% | 6.395% | 4.426% | 0% | 8.447% |  | 4.016% |  | 1.553% | 5.766% | 5.238% | 5.212% | 4.016% | 3.355% | 1.52e-14%
| PDM3 |  | 3.502% | 4.502% | 8.137% | 20.5% | -10.4% |  | -5.767% |  | 0.7369% | -11.95% | -8.952% | 45.21% | -5.767% | 12.77% | -7.312%
| Z-Score (Absolute) |  | 0.19 | 0.26 | 0.52 | 1.38 | 0.78 | 315.54 | 0.45 | 3.79 | 0 | 0.88 | 0.67 | 3.1 | 0.45 | 0.84 | 0.56
| Analytical results for Tm in OREAS 254c (Indicative Value 0.11 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.119 | 0.099585 | 0.1 | [16 empty] | 1
| 1 | 2 | 0.123 | 0.10205 | 0.1 | [16 empty] | 7
| 1 | 3 | 0.111 | 0.10521 | 0.1 | [16 empty] | 16
| 1 | 4 | 0.123 | 0.10797 | 0.1 | [16 empty] | 0.10728
| 1 | 5 | 0.113 | 0.10384 | 0.1 | [16 empty] | 13
| 1 | 6 | 0.116 | 0.10731 | 0.1
| Mean |  | 0.1175 | 0.10433 | 0.1
| Median |  | 0.1175 | 0.10453 | 0.1
| Std Dev. |  | 0.0050498 | 0.0031899 | 1.5202e-17
| Rel.Std.Dev. |  | 4.298% | 3.058% | 1.52e-14%
| PDM3 |  | 9.53% | -2.748% | -6.783%
| Z-Score (Absolute) |  | 2.05 | 0 | 0.67
| Analytical results for U in OREAS 254c (Certified Value 0.92 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.9 | 0.90725 | 0.91 | 0.9 | 0.84 | 0.92 | 0.9 | 0.977 | 0.77 | 0.91493 | 1.04 | 0.93 | 0.99 | 0.83 | 1.15 |  |  |  |  | 1
| 1 | 2 | 0.94 | 0.91872 | 0.93 | 0.9 | 0.85 | 0.89 | 0.9 | 0.963 | 0.76 | 0.90752 | 1.04 | 0.89 | 1.03 | 0.85 | 1.13 |  |  |  |  | 30
| 1 | 3 | 0.91 | 0.92382 | 0.92 | 0.9 | 0.87 | 0.9 | 0.9 | 0.986 | 0.78 | 0.91605 | 1.05 | 0.88 | 0.98 | 0.85 | 1.2 |  |  |  |  | 16
| 1 | 4 | 0.93 | 0.93248 | 0.92 | 0.9 | 0.89 | 0.92 | 0.9 | 0.963 | 0.82 | 0.90649 | 1.01 | 0.89 | 1.01 | 0.84 | 1.17 |  |  |  |  | 0.91529
| 1 | 5 | 0.89 | 0.92087 | 0.87 | 0.9 | 0.87 | 0.9 | 0.9 | 0.949 | 0.87 | 0.92881 | 1.04 | 0.92 | 1.05 | 0.84 | 1.15 |  |  |  |  | 114
| 1 | 6 | 0.9 | 0.9285 | 0.89 | 0.9 | 0.86 | 0.91 | 0.9 | 0.967 | 0.8 | 0.93013 | 1.04 | 0.86 | 1.01 | 0.86 | 1.15
| Mean |  | 0.91167 | 0.92194 | 0.90667 | 0.9 | 0.86333 | 0.90667 | 0.9 | 0.9675 | 0.8 | 0.91732 | 1.0367 | 0.895 | 1.0117 | 0.845 | 1.1583
| Median |  | 0.905 | 0.92234 | 0.915 | 0.9 | 0.865 | 0.905 | 0.9 | 0.965 | 0.79 | 0.91549 | 1.04 | 0.89 | 1.01 | 0.845 | 1.15
| Std Dev. |  | 0.019408 | 0.0087748 | 0.022509 | 0 | 0.017512 | 0.012111 | 0 | 0.012771 | 0.040497 | 0.010167 | 0.013663 | 0.025884 | 0.025626 | 0.010488 | 0.024014
| Rel.Std.Dev. |  | 2.129% | 0.9518% | 2.483% | 0% | 2.028% | 1.336% | 0% | 1.32% | 5.062% | 1.108% | 1.318% | 2.892% | 2.533% | 1.241% | 2.073%
| PDM3 |  | -0.3954% | 0.7268% | -0.9417% | -1.67% | -5.676% | -0.9417% | -1.67% | 5.705% | -12.6% | 0.2225% | 13.26% | -2.216% | 10.53% | -7.679% | 26.55%
| Z-Score (Absolute) |  | 0 | 0.14 | 0.07 | NA | 0.67 | 0.07 | NA | 0.78 | 1.56 | 0.08 | 1.74 | 0.23 | 1.4 | 0.93 | 3.44
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 254c (Certified Value 34.1 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 35 | 34.589 | 32.482 | 33 | 39 | 31 | 30 | 36.9004 | 34 | 35 | 32 | 35.6663 | 35 | 36 | 34 | 33 |  |  |  | 1
| 1 | 2 | 36 | 34.9028 | 33.047 | 30 | 39 | 32 | 30 | 36.4777 | 33 | 36 | 32 | 34.8072 | 35 | 34 | 35 | 32 |  |  |  | 31
| 1 | 3 | 36 | 34.3587 | 32.432 | 34 | 38 | 33 | 30 | 37.0618 | 34 | 35 | 32 | 35.3479 | 36 | 36 | 32 | 32 |  |  |  | 16
| 1 | 4 | 35 | 35.3474 | 32.611 | 33 | 37 | 32 | 30 | 35.713 | 34 | 34 | 33 | 34.7715 | 36 | 35 | 35 | 32 |  |  |  | 34.0941
| 1 | 5 | 35 | 34.58 | 32.725 | 30 | 37 | 32 | 30 | 36.8361 | 33 | 37 | 33 | 36.4365 | 37 | 36 | 34 | 32 |  |  |  | 115
| 1 | 6 | 34 | 35.0139 | 33.313 | 32 | 38 | 31 | 29 | 35.93 | 34 | 35 | 34 | 35.6868 | 36 | 35 | 35 | 32
| Mean |  | 35.1667 | 34.7986 | 32.7683 | 32 | 38 | 31.8333 | 29.8333 | 36.4865 | 33.6667 | 35.3333 | 32.6667 | 35.4527 | 35.8333 | 35.3333 | 34.1667 | 32.1667
| Median |  | 35 | 34.7459 | 32.668 | 32.5 | 38 | 32 | 30 | 36.6569 | 34 | 35 | 32.5 | 35.5071 | 36 | 35.5 | 34.5 | 32
| Std Dev. |  | 0.75277 | 0.35877 | 0.34527 | 1.6733 | 0.89443 | 0.75277 | 0.40825 | 0.55362 | 0.5164 | 1.0328 | 0.8165 | 0.62609 | 0.75277 | 0.8165 | 1.169 | 0.40825
| Rel.Std.Dev. |  | 2.141% | 1.031% | 1.054% | 5.229% | 2.354% | 2.365% | 1.368% | 1.517% | 1.534% | 2.923% | 2.499% | 1.766% | 2.101% | 2.311% | 3.422% | 1.269%
| PDM3 |  | 3.146% | 2.066% | -3.889% | -6.142% | 11.46% | -6.631% | -12.5% | 7.017% | -1.254% | 3.635% | -4.187% | 3.985% | 5.101% | 3.635% | 0.2127% | -5.653%
| Z-Score (Absolute) |  | 0.3 | 0.14 | 0.75 | 1.09 | 1.55 | 1.17 | 2.05 | 0.88 | 0.36 | 0.37 | 0.8 | 0.43 | 0.59 | 0.37 | 0.14 | 1.02
| Analytical results for W in OREAS 254c (Certified Value 0.56 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.53 | < 1 | 1 | 0.3 | 0.4 | 0.3 | < 5 | 0.58 | 0.5 | 0.56 | 0.51235 | 0.4 | 0.58 | 0.71 | 0.44 |  |  |  |  | 1
| 1 | 2 | 0.56 | < 1 | 1 | 0.3 | 0.4 | 0.31 | < 5 | 0.59 | 0.5 | 0.6 | 0.47232 | 0.6 | 0.58 | 0.69 | 0.44 |  |  |  |  | 32
| 1 | 3 | 0.55 | < 1 | < 1 | 0.3 | 0.5 | 0.31 | < 5 | 0.57 | 0.4 | 0.57 | 0.54495 | 0.5 | 0.53 | 0.62 | 0.51 |  |  |  |  | 16
| 1 | 4 | 0.57 | < 1 | 1 | 0.4 | 0.4 | 0.31 | < 5 | 0.61 | 0.4 | 0.59 | 0.54049 | 0.4 | 0.55 | 0.68 | 0.46 |  |  |  |  | 0.56373
| 1 | 5 | 0.53 | < 1 | < 1 | 0.4 | 0.4 | 0.31 | < 5 | 0.66 | 0.4 | 0.58 | 0.52 | 0.6 | 0.58 | 0.66 | 0.48 |  |  |  |  | 116
| 1 | 6 | 0.56 | < 1 | 1 | 0.3 | 0.4 | 0.3 | < 5 | 0.59 | 0.4 | 0.6 | 0.54852 | 0.5 | 0.56 | 0.66 | 0.48
| Mean |  | 0.55 |  | 1 | 0.33333 | 0.41667 | 0.30667 |  | 0.6 | 0.43333 | 0.58333 | 0.52311 | 0.5 | 0.56333 | 0.67 | 0.46833
| Median |  | 0.555 |  | 1 | 0.3 | 0.4 | 0.31 |  | 0.59 | 0.4 | 0.585 | 0.53025 | 0.5 | 0.57 | 0.67 | 0.47
| Std Dev. |  | 0.016733 |  | 0 | 0.05164 | 0.040825 | 0.005164 |  | 0.032249 | 0.05164 | 0.01633 | 0.028739 | 0.089443 | 0.020656 | 0.030984 | 0.027142
| Rel.Std.Dev. |  | 3.042% |  | 0% | 15.49% | 9.798% | 1.684% |  | 5.375% | 11.92% | 2.799% | 5.494% | 17.89% | 3.667% | 4.624% | 5.795%
| PDM3 |  | -2.435% |  | 77.39% | -40.87% | -26.09% | -45.6% |  | 6.434% | -23.13% | 3.478% | -7.206% | -11.3% | -0.07028% | 18.85% | -16.92%
| Z-Score (Absolute) |  | 0.09 | 0.76 | NA | NA | NA | 3.37 | 26.21 | 0.58 | NA | 0.36 | 0.45 | NA | 0.09 | 1.53 | 1.19
Results from laboratories 1.07, 1.09, 1.13 and 1.17 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 254c (Certified Value 8.9 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 8.93 | 8.9482 | 10 | 8.8 | 8.46 | 8.03 | 9.12 | 8.76 | 9.02 | 9.0109 | 8.8 | 9.45 | 9.64 | 9.05 | 10.91 |  |  |  |  | 1
| 1 | 2 | 9.08 | 9.0751 | 9 | 8.76 | 8.74 | 8.13 | 9.0433 | 8.75 | 9.05 | 8.8383 | 8.8 | 9.07 | 9.93 | 9.03 | 10.94 |  |  |  |  | 33
| 1 | 3 | 9.06 | 8.9747 | 10 | 8.6 | 8.96 | 8.39 | 9.3533 | 8.66 | 9.01 | 8.9222 | 9.1 | 9.39 | 9.55 | 8.72 | 11.02 |  |  |  |  | 16
| 1 | 4 | 9.41 | 9.1925 | 10 | 8.33 | 8.59 | 8.4 | 8.83 | 8.82 | 9.82 | 8.7967 | 9 | 9.01 | 10.05 | 8.97 | 10.59 |  |  |  |  | 8.9008
| 1 | 5 | 8.95 | 9.1175 | 9 | 8.47 | 8.58 | 8.32 | 9.31 | 8.73 | 9.59 | 9.0685 | 9.1 | 9.45 | 10.65 | 8.86 | 10.55 |  |  |  |  | 117
| 1 | 6 | 9.05 | 9.1033 | 9 | 8.42 | 8.63 | 8.1 | 8.7633 | 8.74 | 9.41 | 9.1466 | 9.1 | 9.28 | 10.35 | 8.91 | 10.39
| Mean |  | 9.08 | 9.0686 | 9.5 | 8.5633 | 8.66 | 8.2283 | 9.07 | 8.7433 | 9.3167 | 8.9639 | 8.9833 | 9.275 | 10.0283 | 8.9233 | 10.7333
| Median |  | 9.055 | 9.0892 | 9.5 | 8.535 | 8.61 | 8.225 | 9.0817 | 8.745 | 9.23 | 8.9665 | 9.05 | 9.335 | 9.99 | 8.94 | 10.75
| Std Dev. |  | 0.17297 | 0.091964 | 0.54772 | 0.1896 | 0.17239 | 0.16092 | 0.24196 | 0.05164 | 0.34349 | 0.13564 | 0.1472 | 0.19326 | 0.41945 | 0.12258 | 0.25618
| Rel.Std.Dev. |  | 1.905% | 1.014% | 5.766% | 2.214% | 1.991% | 1.956% | 2.668% | 0.5906% | 3.687% | 1.513% | 1.639% | 2.084% | 4.183% | 1.374% | 2.387%
| PDM3 |  | 2.013% | 1.885% | 6.732% | -3.792% | -2.706% | -7.555% | 1.901% | -1.769% | 4.672% | 0.7083% | 0.9271% | 4.204% | 12.67% | 0.253% | 20.59%
| Z-Score (Absolute) |  | 0.14 | 0.11 | NA | 1.17 | 0.93 | 2.02 | 0.11 | 0.72 | 0.74 | 0.16 | 0.11 | 0.63 | 2.54 | 0.26 | 4.33
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 254c (Certified Value 0.64 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.17 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.25g | [14 empty] | 3
| 1 | 1 | 0.664 | 0.63006 | 0.6 | 0.6 | 0.7 | [14 empty] | 1
| 1 | 2 | 0.658 | 0.59477 | 0.7 | 0.6 | 0.7 | [14 empty] | 34
| 1 | 3 | 0.658 | 0.59652 | 0.7 | 0.6 | 0.7 | [14 empty] | 16
| 1 | 4 | 0.641 | 0.63176 | 0.6 | 0.6 | 0.7 | [14 empty] | 0.64125
| 1 | 5 | 0.644 | 0.62358 | 0.6 | 0.6 | 0.7 | [14 empty] | 118
| 1 | 6 | 0.653 | 0.64272 | 0.6 | 0.6 | 0.7
| Mean |  | 0.653 | 0.6199 | 0.63333 | 0.6 | 0.7
| Median |  | 0.6555 | 0.62682 | 0.6 | 0.6 | 0.7
| Std Dev. |  | 0.0088994 | 0.019781 | 0.05164 | 0 | 1.2162e-16
| Rel.Std.Dev. |  | 1.363% | 3.191% | 8.154% | 0% | 1.737e-14%
| PDM3 |  | 1.833% | -3.329% | -1.234% | -6.432% | 9.162%
| Z-Score (Absolute) |  | 0.67 | 0.46 | 0 | 1.14 | 2.29
| Analytical results for Zn in OREAS 254c (Certified Value 46.6 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 2
| 1 | 1 | 46 | 54 | 46.6526 | 37.061 | 43 | 51 | 54.8 | 43 | 41.1467 | 47 | 47 | 43.958 | 44 | 48.0213 | 52.9 | 49 | 46 | 45 | 40 | 1
| 1 | 2 | 47 | 55.998 | 47.0465 | 39.097 | 40 | 50 | 53.4 | 46 | 41.1233 | 46 | 47 | 43.951 | 44 | 46.6221 | 52.1 | 47 | 47 | 43 | 40 | 35
| 1 | 3 | 47 | 53.3304 | 46.5393 | 37.33 | 48 | 50 | 52.6 | 45 | 42.4933 | 46 | 46 | 44.575 | 44 | 46.131 | 54 | 49 | 44 | 43 | 40 | 16
| 1 | 4 | 48 | 52.3692 | 47.6371 | 37.227 | 44 | 50 | 53.9 | 44 | 40.0833 | 47 | 44 | 44.273 | 46 | 46.6272 | 53.1 | 48 | 46 | 43 | 39 | 46.646
| 1 | 5 | 46 | 54.2592 | 46.9744 | 37.731 | 45 | 50 | 52.8 | 44 | 42.4233 | 46 | 48 | 43.092 | 45 | 47.8896 | 54.5 | 49 | 45 | 43 | 38 | 119
| 1 | 6 | 47 | 53.5032 | 47.5297 | 38.927 | 42 | 50 | 52.2 | 43 | 39.73 | 47 | 45 | 44.429 | 47 | 48.9645 | 53.1 | 48 | 48 | 43 | 39
| Mean |  | 46.8333 | 53.91 | 47.0633 | 37.8955 | 43.6667 | 50.1667 | 53.2833 | 44.1667 | 41.1667 | 46.5 | 46.1667 | 44.0463 | 45 | 47.3759 | 53.2833 | 48.3333 | 46 | 43.3333 | 39.3333
| Median |  | 47 | 53.7516 | 47.0104 | 37.5305 | 43.5 | 50 | 53.1 | 44 | 41.135 | 46.5 | 46.5 | 44.1155 | 44.5 | 47.2584 | 53.1 | 48.5 | 46 | 43 | 39.5
| Std Dev. |  | 0.75277 | 1.2137 | 0.44679 | 0.89417 | 2.7325 | 0.40825 | 0.95586 | 1.169 | 1.1471 | 0.54772 | 1.472 | 0.53008 | 1.2649 | 1.0847 | 0.84951 | 0.8165 | 1.4142 | 0.8165 | 0.8165
| Rel.Std.Dev. |  | 1.607% | 2.251% | 0.9493% | 2.36% | 6.258% | 0.8138% | 1.794% | 2.647% | 2.786% | 1.178% | 3.188% | 1.203% | 2.811% | 2.29% | 1.594% | 1.689% | 3.074% | 1.884% | 2.076%
| PDM3 |  | 0.4015% | 15.57% | 0.8944% | -18.76% | -6.387% | 7.548% | 14.23% | -5.315% | -11.75% | -0.3131% | -1.028% | -5.573% | -3.529% | 1.565% | 14.23% | 3.617% | -1.385% | -7.102% | -15.68%
| Z-Score (Absolute) |  | 0.21 | 2.41 | 0.28 | 2.57 | 0.78 | 1.24 | 2.21 | 0.62 | 1.56 | 0.1 | 0 | 0.66 | 0.36 | 0.38 | 2.21 | 0.67 | 0.05 | 0.88 | 2.13
| Analytical results for Zr in OREAS 254c (Certified Value 34.4 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 13.8 | 44.7696 | 90 | 45.4 | 6.4 | 33 | 34.5692 | 31.4 | 31.3 | 59.1313 | 49.2 | 25.5 | 40.2 | 22.2 | 52.2 |  |  |  |  | 1
| 1 | 2 | 15.5 | 45.2972 | 86 | 43.9 | 5.8 | 33 | 34.6787 | 34.1 | 38 | 49.3991 | 46.7 | 32 | 37.1 | 22.6 | 52.7 |  |  |  |  | 36
| 1 | 3 | 16.9 | 43.7207 | 93 | 42.4 | 6.4 | 32 | 36.0883 | 34.2 | 31.8 | 55.173 | 46.4 | 23.3 | 35.9 | 28.3 | 50.4 |  |  |  |  | 16
| 1 | 4 | 16.2 | 50.6198 | 88 | 41.9 | 7.1 | 31 | 34.0714 | 35.7 | 25.5 | 51.7436 | 47.3 | 24.7 | 43.2 | 22.7 | 51 |  |  |  |  | 34.4295
| 1 | 5 | 14 | 44.5942 | 170 | 43.4 | 5.4 | 34 | 36.3003 | 34.8 | 29.2 | 59.9206 | 47.5 | 27.5 | 37 | 20.2 | 49.3 |  |  |  |  | 120
| 1 | 6 | 15.3 | 48.4773 | 88 | 41.3 | 7.1 | 31 | 33.753 | 32.9 | 28.4 | 49.2706 | 48.7 | 26.3 | 37.9 | 23.7 | 50.9
| Mean |  | 15.2833 | 46.2465 | 102.5 | 43.05 | 6.3667 | 32.3333 | 34.9101 | 33.85 | 30.7 | 54.1064 | 47.6333 | 26.55 | 38.55 | 23.2833 | 51.0833
| Median |  | 15.4 | 45.0334 | 89 | 42.9 | 6.4 | 32.5 | 34.624 | 34.15 | 30.25 | 53.4583 | 47.4 | 25.9 | 37.5 | 22.65 | 50.95
| Std Dev. |  | 1.2123 | 2.6942 | 33.1527 | 1.495 | 0.68313 | 1.2111 | 1.0519 | 1.511 | 4.2294 | 4.7185 | 1.1057 | 3.0251 | 2.6943 | 2.714 | 1.2287
| Rel.Std.Dev. |  | 7.932% | 5.826% | 32.34% | 3.473% | 10.73% | 3.746% | 3.013% | 4.464% | 13.78% | 8.721% | 2.321% | 11.39% | 6.989% | 11.66% | 2.405%
| PDM3 |  | -55.61% | 34.32% | 197.7% | 25.04% | -81.51% | -6.088% | 1.396% | -1.683% | -10.83% | 57.15% | 38.35% | -22.89% | 11.97% | -32.37% | 48.37%
| Z-Score (Absolute) |  | 1.17 | 0.67 | 4.02 | 0.48 | 1.7 | 0.15 | 0 | 0.06 | 0.25 | 1.14 | 0.76 | 0.5 | 0.22 | 0.69 | 0.96
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 254c (Indicative Value 9.85 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 9.83 | 1
| 1 | 2 | 9.86 | 11
| Mean |  | 9.845 | 16
| Median |  | 9.845 | 9.845
| Std Dev. |  | 0.021213 | 17
| Rel.Std.Dev. |  | 0.2155%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 254c (Indicative Value 1.99 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.99 | 1
| 1 | 2 | 1.99 | 12
| Mean |  | 1.99 | 16
| Median |  | 1.99 | 1.99
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 254c (Indicative Value 5.73 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.74 | 1
| 1 | 2 | 5.72 | 13
| Mean |  | 5.73 | 16
| Median |  | 5.73 | 5.73
| Std Dev. |  | 0.014142 | 19
| Rel.Std.Dev. |  | 0.2468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 254c (Indicative Value 1.53 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.53 | 14
| Mean |  | 1.53 | 16
| Median |  | 1.53 | 1.53
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 254c (Indicative Value 2.57 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.57 | 1
| 1 | 2 | 2.57 | 11
| Mean |  | 2.57 | 16
| Median |  | 2.57 | 2.57
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 254c (Indicative Value 0.059 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.059 | 1
| 1 | 2 | 0.059 | 12
| Mean |  | 0.059 | 16
| Median |  | 0.059 | 0.059
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 254c (Indicative Value 1.31 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.31 | 1
| 1 | 2 | 1.31 | 13
| Mean |  | 1.31 | 16
| Median |  | 1.31 | 1.31
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 254c (Indicative Value 0.201 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.201 | 1
| 1 | 2 | 0.2 | 14
| Mean |  | 0.2005 | 16
| Median |  | 0.2005 | 0.2005
| Std Dev. |  | 0.00070711 | 20
| Rel.Std.Dev. |  | 0.3527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 254c (Indicative Value 73.26 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 73.26 | 1
| 1 | 2 | 73.26 | 11
| Mean |  | 73.26 | 16
| Median |  | 73.26 | 73.26
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 254c (Indicative Value 0.054 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.054 | 1
| 1 | 2 | 0.054 | 12
| Mean |  | 0.054 | 16
| Median |  | 0.054 | 0.054
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 254c (Indicative Value 0.931 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.931 | 1
| 1 | 2 | 0.93 | 13
| Mean |  | 0.9305 | 16
| Median |  | 0.9305 | 0.9305
| Std Dev. |  | 0.00070711 | 19
| Rel.Std.Dev. |  | 0.07599%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 254c (Indicative Value 2.32 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.3 | 1
| 1 | 2 | 2.33 | 16
| Mean |  | 2.315 | 16
| Median |  | 2.315 | 2.315
| Std Dev. |  | 0.021213 | 22
| Rel.Std.Dev. |  | 0.9163%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 254c (Indicative Value 0.085 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.09 | 1
| 1 | 2 | 0.08 | 18
| Mean |  | 0.085 | 16
| Median |  | 0.085 | 0.085
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 254c (Indicative Value < 0.01 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 18
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 24
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 254c (Indicative Value 1 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1 | 20
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 254c (Indicative Value 17.1 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.2 | 1
| 1 | 2 | 17 | 21
| Mean |  | 17.1 | 16
| Median |  | 17.1 | 17.1
| Std Dev. |  | 0.14142 | 27
| Rel.Std.Dev. |  | 0.827%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 254c (Indicative Value 343 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 339 | 1
| 1 | 2 | 346 | 22
| Mean |  | 342.5 | 16
| Median |  | 342.5 | 342.5
| Std Dev. |  | 4.9497 | 28
| Rel.Std.Dev. |  | 1.445%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 254c (Indicative Value 2.1 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 1.8 | 23
| Mean |  | 2.1 | 16
| Median |  | 2.1 | 2.1
| Std Dev. |  | 0.42426 | 29
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 254c (Indicative Value 0.35 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.34 | 24
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.014142 | 30
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 254c (Indicative Value 0.075 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | < 0.1 | 25
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.075
| Std Dev. |  |  | 31
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 254c (Indicative Value 68 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 68 | 1
| 1 | 2 | 67.3 | 26
| Mean |  | 67.65 | 16
| Median |  | 67.65 | 67.65
| Std Dev. |  | 0.49497 | 32
| Rel.Std.Dev. |  | 0.7317%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 254c (Indicative Value 20.5 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.6 | 1
| 1 | 2 | 20.3 | 27
| Mean |  | 20.45 | 16
| Median |  | 20.45 | 20.45
| Std Dev. |  | 0.21213 | 33
| Rel.Std.Dev. |  | 1.037%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 254c (Indicative Value 128 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 128 | 1
| 1 | 2 | 128 | 28
| Mean |  | 128 | 16
| Median |  | 128 | 128
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 254c (Indicative Value 2.45 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.44 | 1
| 1 | 2 | 2.46 | 29
| Mean |  | 2.45 | 16
| Median |  | 2.45 | 2.45
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 0.5772%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 254c (Indicative Value 33 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32 | 1
| 1 | 2 | 34 | 30
| Mean |  | 33 | 16
| Median |  | 33 | 33
| Std Dev. |  | 1.4142 | 36
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 254c (Indicative Value 4.01 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 4.02 | 4
| Mean |  | 4.01 | 16
| Median |  | 4.01 | 4.01
| Std Dev. |  | 0.014142 | 37
| Rel.Std.Dev. |  | 0.3527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 254c (Indicative Value 1.97 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.99 | 1
| 1 | 2 | 1.95 | 5
| Mean |  | 1.97 | 16
| Median |  | 1.97 | 1.97
| Std Dev. |  | 0.028284 | 38
| Rel.Std.Dev. |  | 1.436%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 254c (Indicative Value 1.6 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.56 | 1
| 1 | 2 | 1.64 | 6
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0.056569 | 39
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 254c (Indicative Value 14.6 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.4 | 1
| 1 | 2 | 14.7 | 34
| Mean |  | 14.55 | 16
| Median |  | 14.55 | 14.55
| Std Dev. |  | 0.21213 | 40
| Rel.Std.Dev. |  | 1.458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 254c (Indicative Value 5.08 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.15 | 1
| 1 | 2 | 5.01 | 7
| Mean |  | 5.08 | 16
| Median |  | 5.08 | 5.08
| Std Dev. |  | 0.098995 | 41
| Rel.Std.Dev. |  | 1.949%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 254c (Indicative Value 1.05 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1.1 | 1
| Mean |  | 1.05 | 16
| Median |  | 1.05 | 1.05
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 254c (Indicative Value 6.1 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.13 | 1
| 1 | 2 | 6.06 | 8
| Mean |  | 6.095 | 16
| Median |  | 6.095 | 6.095
| Std Dev. |  | 0.049497 | 26
| Rel.Std.Dev. |  | 0.8121%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 254c (Indicative Value 0.75 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.73 | 9
| Mean |  | 0.745 | 16
| Median |  | 0.745 | 0.745
| Std Dev. |  | 0.021213 | 27
| Rel.Std.Dev. |  | 2.847%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 254c (Indicative Value 0.038 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.05 | 1
| 1 | 2 | 0.05 | 22
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.0375
| Std Dev. |  |  | 28
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 254c (Indicative Value 38.6 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.6 | 1
| 1 | 2 | 38.5 | 23
| Mean |  | 38.55 | 16
| Median |  | 38.55 | 38.55
| Std Dev. |  | 0.070711 | 29
| Rel.Std.Dev. |  | 0.1834%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 254c (Indicative Value 0.22 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.22 | 4
| Mean |  | 0.22 | 16
| Median |  | 0.22 | 0.22
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 254c (Indicative Value 0.047 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.047 | 1
| 1 | 2 | 0.047 | 25
| Mean |  | 0.047 | 16
| Median |  | 0.047 | 0.047
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 254c (Indicative Value 3.8 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 3.8 | 26
| Mean |  | 3.8 | 16
| Median |  | 3.8 | 3.8
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 254c (Indicative Value 32.2 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32.2 | 1
| 1 | 2 | 32.2 | 5
| Mean |  | 32.2 | 16
| Median |  | 32.2 | 32.2
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 254c (Indicative Value 32.7 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32.9 | 1
| 1 | 2 | 32.5 | 6
| Mean |  | 32.7 | 16
| Median |  | 32.7 | 32.7
| Std Dev. |  | 0.28284 | 34
| Rel.Std.Dev. |  | 0.865%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 254c (Indicative Value 90 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 90 | 29
| Mean |  | 90 | 16
| Median |  | 90 | 90
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 254c (Indicative Value 11 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11 | 1
| 1 | 2 | 11 | 30
| Mean |  | 11 | 16
| Median |  | 11 | 11
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 254c (Indicative Value 8.86 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.89 | 1
| 1 | 2 | 8.82 | 8
| Mean |  | 8.855 | 16
| Median |  | 8.855 | 8.855
| Std Dev. |  | 0.049497 | 37
| Rel.Std.Dev. |  | 0.559%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 254c (Indicative Value 62 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 62.8 | 1
| 1 | 2 | 61.8 | 32
| Mean |  | 62.3 | 16
| Median |  | 62.3 | 62.3
| Std Dev. |  | 0.70711 | 38
| Rel.Std.Dev. |  | 1.135%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 254c (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 33
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 39
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 254c (Indicative Value 1.25 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.3 | 34
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 254c (Indicative Value 9.35 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.3 | 1
| 1 | 2 | 9.4 | 35
| Mean |  | 9.35 | 16
| Median |  | 9.35 | 9.35
| Std Dev. |  | 0.070711 | 41
| Rel.Std.Dev. |  | 0.7563%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 254c (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 1
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 42
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 254c (Indicative Value 6.17 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.04 | 1
| 1 | 2 | 6.3 | 5
| Mean |  | 6.17 | 16
| Median |  | 6.17 | 6.17
| Std Dev. |  | 0.18385 | 26
| Rel.Std.Dev. |  | 2.98%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 254c (Indicative Value 5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5 | 1
| 1 | 2 | 5 | 21
| Mean |  | 5 | 16
| Median |  | 5 | 5
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 254c (Indicative Value 278 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 277 | 1
| 1 | 2 | 278 | 22
| Mean |  | 277.5 | 16
| Median |  | 277.5 | 277.5
| Std Dev. |  | 0.70711 | 28
| Rel.Std.Dev. |  | 0.2548%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 254c (Indicative Value 2.11 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.15 | 1
| 1 | 2 | 2.07 | 23
| Mean |  | 2.11 | 16
| Median |  | 2.11 | 2.11
| Std Dev. |  | 0.056569 | 29
| Rel.Std.Dev. |  | 2.681%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 254c (Indicative Value 0.75 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.74 | 24
| Mean |  | 0.75 | 16
| Median |  | 0.75 | 0.75
| Std Dev. |  | 0.014142 | 30
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 254c (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 2
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 31
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 254c (Indicative Value 8.76 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.8 | 1
| 1 | 2 | 8.71 | 26
| Mean |  | 8.755 | 16
| Median |  | 8.755 | 8.755
| Std Dev. |  | 0.06364 | 32
| Rel.Std.Dev. |  | 0.7269%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 254c (Indicative Value 0.563 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.562 | 1
| 1 | 2 | 0.563 | 27
| Mean |  | 0.5625 | 16
| Median |  | 0.5625 | 0.5625
| Std Dev. |  | 0.00070711 | 33
| Rel.Std.Dev. |  | 0.1257%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 254c (Indicative Value 0.15 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | 0.2 | 28
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 34
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 254c (Indicative Value 0.27 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.26 | 1
| 1 | 2 | 0.27 | 7
| Mean |  | 0.265 | 16
| Median |  | 0.265 | 0.265
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 254c (Indicative Value 1.79 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.75 | 1
| 1 | 2 | 1.83 | 30
| Mean |  | 1.79 | 16
| Median |  | 1.79 | 1.79
| Std Dev. |  | 0.056569 | 36
| Rel.Std.Dev. |  | 3.16%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 254c (Indicative Value 73 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72.9 | 1
| 1 | 2 | 72.9 | 31
| Mean |  | 72.9 | 16
| Median |  | 72.9 | 72.9
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 254c (Indicative Value 3.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 3.5 | 32
| Mean |  | 3.75 | 16
| Median |  | 3.75 | 3.75
| Std Dev. |  | 0.35355 | 38
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 254c (Indicative Value 19.1 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.2 | 1
| 1 | 2 | 19 | 33
| Mean |  | 19.1 | 16
| Median |  | 19.1 | 19.1
| Std Dev. |  | 0.14142 | 39
| Rel.Std.Dev. |  | 0.7404%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 254c (Indicative Value 1.75 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.7 | 1
| 1 | 2 | 1.79 | 34
| Mean |  | 1.745 | 16
| Median |  | 1.745 | 1.745
| Std Dev. |  | 0.06364 | 40
| Rel.Std.Dev. |  | 3.647%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 254c (Indicative Value 75 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75 | 1
| 1 | 2 | 75 | 35
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 254c (Indicative Value 244 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 245 | 1
| 1 | 2 | 243 | 36
| Mean |  | 244 | 16
| Median |  | 244 | 244
| Std Dev. |  | 1.4142 | 42
| Rel.Std.Dev. |  | 0.5796%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.