## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 250c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.3125 | 0.3061 | 0.3189 | 0.3102 | 0.3148
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 0.31298 | 0.30462 | 0.32133 | 0.31045 | 0.3155
Cyanide Leach
Au, Gold (ppm) | 0.29515 | 0.28697 | 0.30334 | 0.29431 | 0.296
PhotonAssay (recommended gross mass* 505-535 g)
Au, Gold (ppm) | 0.31092 | 0.3022 | 0.31965 | 0.31025 | 0.3116
4-Acid Digestion
Ag, Silver (ppm) | < 0.1 | IND | IND | IND | IND
Al, Aluminium (wt.%) | 5.176 | 5.0403 | 5.3117 | 5.0983 | 5.2538
As, Arsenic (ppm) | 54.0987 | 51.4842 | 56.7132 | 51.9764 | 56.221
Ba, Barium (ppm) | 354.248 | 342.3428 | 366.1533 | 345.5714 | 362.9247
Be, Beryllium (ppm) | 1.9732 | 1.8931 | 2.0533 | 1.8914 | 2.0549
Bi, Bismuth (ppm) | 0.29075 | 0.26323 | 0.31826 | 0.27428 | 0.30721
Ca, Calcium (wt.%) | 1.7157 | 1.6574 | 1.7741 | 1.6774 | 1.754
Cd, Cadmium (ppm) | 0.074604 | 0.044181 | 0.10503 | 0.064495 | 0.084713
Ce, Cerium (ppm) | 70.4375 | 67.0478 | 73.8272 | 67.2433 | 73.6317
Co, Cobalt (ppm) | 20.5756 | 19.4939 | 21.6574 | 20.0365 | 21.1147
Cr, Chromium (ppm) | 100.2348 | 94.7049 | 105.7646 | 96.3667 | 104.1028
Cs, Caesium (ppm) | 2.6165 | 2.4841 | 2.749 | 2.5239 | 2.7091
Cu, Copper (ppm) | 24.1278 | 22.6828 | 25.5728 | 23.1706 | 25.0849
Dy, Dysprosium (ppm) | 3.6654 | 3.3402 | 3.9907 | 3.4389 | 3.892
Er, Erbium (ppm) | 1.5937 | 1.4404 | 1.747 | 1.497 | 1.6905
Eu, Europium (ppm) | 1.6023 | 1.4529 | 1.7517 | 1.5109 | 1.6937
Fe, Iron (wt.%) | 4.0366 | 3.9026 | 4.1706 | 3.9425 | 4.1307
Ga, Gallium (ppm) | 15.3212 | 14.249 | 16.3934 | 14.8391 | 15.8033
Gd, Gadolinium (ppm) | 5.2047 | 4.6724 | 5.737 | 5.0104 | 5.399
Hf, Hafnium (ppm) | 4.3446 | 4.0788 | 4.6104 | 4.106 | 4.5831
Ho, Holmium (ppm) | 0.62908 | 0.57324 | 0.68491 | 0.59257 | 0.66558
In, Indium (ppm) | 0.058553 | 0.04895 | 0.068157 | 0.052688 | 0.064418
K, Potassium (wt.%) | 1.3185 | 1.2727 | 1.3644 | 1.2879 | 1.3491
La, Lanthanum (ppm) | 37.9907 | 35.8002 | 40.1812 | 36.8228 | 39.1586
Li, Lithium (ppm) | 21.7332 | 20.7044 | 22.7621 | 20.987 | 22.4795
Lu, Lutetium (ppm) | 0.17769 | 0.1418 | 0.21359 | IND | IND
Mg, Magnesium (wt.%) | 1.5542 | 1.5 | 1.6083 | 1.5167 | 1.5916
Mn, Manganese (wt.%) | 0.049236 | 0.047496 | 0.050977 | 0.048178 | 0.050295
Mo, Molybdenum (ppm) | 1.9642 | 1.8115 | 2.1169 | 1.8952 | 2.0333
Na, Sodium (wt.%) | 1.0741 | 1.0369 | 1.1112 | 1.0511 | 1.097
Nb, Niobium (ppm) | 31.7716 | 30.0814 | 33.4618 | 30.5726 | 32.9706
Nd, Neodymium (ppm) | 32.0964 | 29.0327 | 35.1602 | 31.1918 | 33.0011
Ni, Nickel (ppm) | 83.1882 | 79.7403 | 86.6362 | 81.0758 | 85.3007
P, Phosphorus (wt.%) | 0.095295 | 0.09181 | 0.098779 | 0.093415 | 0.097175
Pb, Lead (ppm) | 9.9208 | 9.0866 | 10.7549 | 9.4146 | 10.4269
Pr, Praseodymium (ppm) | 8.5113 | 7.7302 | 9.2924 | 8.0967 | 8.9259
Rb, Rubidium (ppm) | 64.7019 | 61.0771 | 68.3267 | 62.9366 | 66.4672
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.027184 | 0.024706 | 0.029662 | 0.026001 | 0.028367
Sb, Antimony (ppm) | 0.88377 | 0.8079 | 0.95964 | 0.84275 | 0.92479
Sc, Scandium (ppm) | 9.8918 | 9.333 | 10.4506 | 9.6347 | 10.1489
Sm, Samarium (ppm) | 6.2719 | 5.8198 | 6.7241 | 5.9722 | 6.5716
Sn, Tin (ppm) | 3.5909 | 3.3659 | 3.8159 | 3.3406 | 3.8412
Sr, Strontium (ppm) | 311.4662 | 299.9595 | 322.9729 | 304.2453 | 318.6871
Ta, Tantalum (ppm) | 2.0748 | 1.9213 | 2.2283 | 1.9456 | 2.2041
Tb, Terbium (ppm) | 0.68798 | 0.6461 | 0.72987 | 0.67432 | 0.70165
Th, Thorium (ppm) | 8.6755 | 8.2309 | 9.1201 | 8.3396 | 9.0114
Ti, Titanium (wt.%) | 0.56692 | 0.54808 | 0.58577 | 0.55674 | 0.57711
Tl, Thallium (ppm) | 0.35326 | 0.3299 | 0.37663 | 0.33464 | 0.37189
Tm, Thulium (ppm) | 0.20127 | 0.18262 | 0.21993 | 0.19285 | 0.2097
U, Uranium (ppm) | 1.6213 | 1.5125 | 1.7301 | 1.5493 | 1.6934
V, Vanadium (ppm) | 69.932 | 67.4764 | 72.3876 | 67.812 | 72.052
W, Tungsten (ppm) | 2.0536 | 1.8969 | 2.2103 | 1.8966 | 2.2105
Y, Yttrium (ppm) | 15.6643 | 14.5995 | 16.7291 | 15.0862 | 16.2424
Yb, Ytterbium (ppm) | 1.2817 | 1.1456 | 1.4178 | 1.1928 | 1.3706
Zn, Zinc (ppm) | 74.9419 | 71.9801 | 77.9038 | 73.3177 | 76.5662
Zr, Zirconium (ppm) | 187.837 | 179.3377 | 196.3363 | 183.8398 | 191.8342
Aqua Regia Digestion
Ag, Silver (ppm) | 0.078574 | 0.066533 | 0.090615 | IND | IND
Al, Aluminium (wt.%) | 1.6421 | 1.5555 | 1.7287 | 1.5899 | 1.6943
As, Arsenic (ppm) | 43.4324 | 42.0126 | 44.8522 | 42.6 | 44.2648
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 94.5276 | 89.2278 | 99.8274 | 91.3381 | 97.7171
Be, Beryllium (ppm) | 1.0776 | 1.0134 | 1.1418 | 1.0429 | 1.1123
Bi, Bismuth (ppm) | 0.2647 | 0.24615 | 0.28325 | 0.24975 | 0.27965
Ca, Calcium (wt.%) | 0.52434 | 0.4932 | 0.55548 | 0.51097 | 0.53772
Cd, Cadmium (ppm) | 0.06726 | 0.047786 | 0.086734 | IND | IND
Ce, Cerium (ppm) | 40.8783 | 39.2184 | 42.5381 | 39.7376 | 42.019
Co, Cobalt (ppm) | 15.9278 | 15.1426 | 16.7129 | 15.0932 | 16.7623
Cr, Chromium (ppm) | 45.8997 | 42.3535 | 49.4458 | 44.4822 | 47.3171
Cs, Caesium (ppm) | 1.1273 | 1.0546 | 1.1999 | 1.0841 | 1.1705
Cu, Copper (ppm) | 18.8888 | 17.6016 | 20.1761 | 18.0966 | 19.6811
Fe, Iron (wt.%) | 3.0079 | 2.8536 | 3.1621 | 2.9465 | 3.0692
Ga, Gallium (ppm) | 5.9673 | 5.4919 | 6.4427 | 5.7849 | 6.1497
Ge, Germanium (ppm) | 0.099667 | 0.076032 | 0.1233 | IND | IND
In, Indium (ppm) | 0.025375 | 0.018689 | 0.032061 | 0.021994 | 0.028756
K, Potassium (wt.%) | 0.33493 | 0.31679 | 0.35308 | 0.32509 | 0.34478
La, Lanthanum (ppm) | 23.162 | 22.3109 | 24.0131 | 22.497 | 23.827
Li, Lithium (ppm) | 7.6061 | 7.1164 | 8.0959 | 7.3159 | 7.8964
Mg, Magnesium (wt.%) | 0.98857 | 0.94215 | 1.035 | 0.96733 | 1.0098
Mn, Manganese (wt.%) | 0.033367 | 0.03192 | 0.034815 | 0.032597 | 0.034137
Mo, Molybdenum (ppm) | 1.3238 | 1.2317 | 1.4158 | 1.2758 | 1.3717
Na, Sodium (wt.%) | 0.35057 | 0.33149 | 0.36965 | 0.33841 | 0.36273
Nb, Niobium (ppm) | 0.58659 | 0.48038 | 0.69279 | 0.5288 | 0.64437
Ni, Nickel (ppm) | 72.6347 | 69.8167 | 75.4526 | 70.5448 | 74.7246
P, Phosphorus (wt.%) | 0.05846 | 0.055663 | 0.061257 | 0.057215 | 0.059705
Pb, Lead (ppm) | 7.0819 | 6.6629 | 7.5009 | 6.8076 | 7.3562
Rb, Rubidium (ppm) | 19.0763 | 18.1017 | 20.0509 | 18.4787 | 19.6739
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.02435 | 0.020453 | 0.028248 | 0.022705 | 0.025996
Sb, Antimony (ppm) | 0.44672 | 0.40383 | 0.48961 | 0.42374 | 0.4697
Sc, Scandium (ppm) | 3.7374 | 3.5095 | 3.9652 | 3.6022 | 3.8725
Sn, Tin (ppm) | 1.2367 | 1.1163 | 1.357 | IND | IND
Sr, Strontium (ppm) | 63.4029 | 61.3024 | 65.5035 | 61.7686 | 65.0373
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.44776 | 0.40146 | 0.49406 | IND | IND
Th, Thorium (ppm) | 5.8503 | 5.4882 | 6.2125 | 5.6586 | 6.042
Ti, Titanium (wt.%) | 0.1622 | 0.15078 | 0.17362 | 0.15724 | 0.16717
Tl, Thallium (ppm) | 0.1356 | 0.12201 | 0.1492 | IND | IND
U, Uranium (ppm) | 0.96523 | 0.89468 | 1.0358 | 0.93989 | 0.99057
V, Vanadium (ppm) | 33.6989 | 32.0967 | 35.3011 | 32.5217 | 34.8761
W, Tungsten (ppm) | 0.38335 | 0.32959 | 0.43711 | 0.32687 | 0.43983
Y, Yttrium (ppm) | 9.1007 | 8.7415 | 9.4599 | 8.9015 | 9.2999
Yb, Ytterbium (ppm) | 0.63153 | 0.55194 | 0.71113 | IND | IND
Zn, Zinc (ppm) | 46.3867 | 43.7579 | 49.0155 | 45.2817 | 47.4917
Zr, Zirconium (ppm) | 37.1153 | 28.6673 | 45.5633 | 35.1565 | 39.0741
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 250c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Ge | ppm | 0.14121 | Se | ppm | 0.75717
Hg | ppm | 0.030057 | Te | ppm | 0.028842
Aqua Regia Digestion
Dy | ppm | 2.2407 | Ho | ppm | 0.35544 | Se | ppm | 0.19292
Er | ppm | 0.86776 | Lu | ppm | 0.090077 | Sm | ppm | 3.9986
Eu | ppm | 0.96937 | Nd | ppm | 22.1759 | Te | ppm | 0.022604
Gd | ppm | 3.3839 | Pd | ppb | < 10 | Tm | ppm | 0.10676
Hf | ppm | 0.67802 | Pr | ppm | 5.7706
Hg | ppm | 0.021667 | Pt | ppb | < 5
Borate Fusion XRF
Al2O3 | wt.% | 10.065 | MgO | wt.% | 2.665 | SiO2 | wt.% | 72.245
CaO | wt.% | 2.35 | MnO | wt.% | 0.064 | SO3 | wt.% | 0.0695
Fe2O3 | wt.% | 5.9 | Na2O | wt.% | 1.46 | TiO2 | wt.% | 0.997
K2O | wt.% | 1.58 | P2O5 | wt.% | 0.2185
Thermogravimetry
LOI1000 | wt.% | 2.155
Infrared Combustion
C | wt.% | 0.075 | S | wt.% | 0.01
Laser Ablation ICP-MS
Ag | ppm | 0.1 | Hf | ppm | 5.985 | Sm | ppm | 6.315
As | ppm | 50.6 | Ho | ppm | 0.77 | Sn | ppm | 4.5
Ba | ppm | 349.5 | In | ppm | 0.0375 | Sr | ppm | 296
Be | ppm | 2 | La | ppm | 38.3 | Ta | ppm | 2.205
Bi | ppm | 0.3 | Lu | ppm | 0.23 | Tb | ppm | 0.77
Cd | ppm | < 0.1 | Mn | wt.% | 0.05085 | Te | ppm | < 0.2
Ce | ppm | 67.85 | Mo | ppm | 1.8 | Th | ppm | 9.09
Co | ppm | 20.5 | Nb | ppm | 32.6 | Ti | wt.% | 0.5985
Cr | ppm | 115 | Nd | ppm | 33.05 | Tl | ppm | 0.15
Cs | ppm | 2.52 | Ni | ppm | 86 | Tm | ppm | 0.285
Cu | ppm | 32 | Pb | ppm | 10.5 | U | ppm | 1.81
Dy | ppm | 4.12 | Pr | ppm | 8.9 | V | ppm | 72.85
Er | ppm | 2.035 | Rb | ppm | 62.85 | W | ppm | 2.5
Eu | ppm | 1.59 | Re | ppm | < 0.01 | Y | ppm | 19.75
Ga | ppm | 14.55 | Sb | ppm | 0.9 | Yb | ppm | 1.745
Gd | ppm | 5.33 | Sc | ppm | 9.55 | Zn | ppm | 75
Ge | ppm | 1.075 | Se | ppm | < 5 | Zr | ppm | 243
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 250c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.3125 | 0.015461 | 0.28158 | 0.34342 | 0.26612 | 0.35888 | 4.948% | 9.895% | 14.84% | 0.29688 | 0.32813
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.31298 | 0.018134 | 0.27671 | 0.34924 | 0.25858 | 0.36738 | 5.794% | 11.59% | 17.38% | 0.29733 | 0.32863
Cyanide Leach
Au, ppm | 0.29515 | 0.016266 | 0.26262 | 0.32769 | 0.24636 | 0.34395 | 5.511% | 11.02% | 16.53% | 0.2804 | 0.30991
PhotonAssay
Au, ppm | 0.31092 | 0.027325 | 0.25627 | 0.36557 | 0.22895 | 0.3929 | 8.788% | 17.58% | 26.36% | 0.29538 | 0.32647
4-Acid Digestion
Ag, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Al, wt.% | 5.176 | 0.10095 | 4.9741 | 5.3779 | 4.8732 | 5.4789 | 1.95% | 3.901% | 5.851% | 4.9172 | 5.4348
As, ppm | 54.0987 | 2.2461 | 49.6064 | 58.591 | 47.3603 | 60.8371 | 4.152% | 8.304% | 12.46% | 51.3938 | 56.8037
Ba, ppm | 354.248 | 6.8352 | 340.5777 | 367.9184 | 333.7426 | 374.7535 | 1.929% | 3.859% | 5.788% | 336.5356 | 371.9604
Be, ppm | 1.9732 | 0.087294 | 1.7986 | 2.1478 | 1.7113 | 2.2351 | 4.424% | 8.848% | 13.27% | 1.8745 | 2.0718
Bi, ppm | 0.29075 | 0.029534 | 0.23168 | 0.34981 | 0.20214 | 0.37935 | 10.16% | 20.32% | 30.47% | 0.27621 | 0.30528
Ca, wt.% | 1.7157 | 0.059922 | 1.5959 | 1.8356 | 1.536 | 1.8955 | 3.492% | 6.985% | 10.48% | 1.6299 | 1.8015
Cd, ppm | 0.074604 | 0.017624 | 0.039355 | 0.10985 | 0.021731 | 0.12748 | 23.62% | 47.25% | 70.87% | 0.070874 | 0.078334
Ce, ppm | 70.4375 | 3.0524 | 64.3328 | 76.5423 | 61.2804 | 79.5946 | 4.333% | 8.667% | 13% | 66.9156 | 73.9594
Co, ppm | 20.5756 | 1.3229 | 17.9299 | 23.2213 | 16.607 | 24.5442 | 6.429% | 12.86% | 19.29% | 19.5468 | 21.6044
Cr, ppm | 100.2348 | 6.1154 | 88.004 | 112.4655 | 81.8887 | 118.5808 | 6.101% | 12.2% | 18.3% | 95.223 | 105.2465
Cs, ppm | 2.6165 | 0.11123 | 2.394 | 2.839 | 2.2828 | 2.9502 | 4.251% | 8.502% | 12.75% | 2.4857 | 2.7473
Cu, ppm | 24.1278 | 2.1823 | 19.7631 | 28.4924 | 17.5808 | 30.6747 | 9.045% | 18.09% | 27.13% | 22.9214 | 25.3342
Dy, ppm | 3.6654 | 0.20416 | 3.2571 | 4.0738 | 3.053 | 4.2779 | 5.57% | 11.14% | 16.71% | 3.4822 | 3.8487
Er, ppm | 1.5937 | 0.1071 | 1.3795 | 1.8079 | 1.2724 | 1.915 | 6.72% | 13.44% | 20.16% | 1.514 | 1.6734
Eu, ppm | 1.6023 | 0.059993 | 1.4823 | 1.7223 | 1.4223 | 1.7823 | 3.744% | 7.488% | 11.23% | 1.5222 | 1.6824
Fe, wt.% | 4.0366 | 0.13643 | 3.7637 | 4.3094 | 3.6273 | 4.4459 | 3.38% | 6.76% | 10.14% | 3.8347 | 4.2384
Ga, ppm | 15.3212 | 1.1231 | 13.075 | 17.5673 | 11.9519 | 18.6904 | 7.33% | 14.66% | 21.99% | 14.5551 | 16.0872
Gd, ppm | 5.2047 | 0.25648 | 4.6917 | 5.7176 | 4.4352 | 5.9741 | 4.928% | 9.856% | 14.78% | 4.9444 | 5.4649
Hf, ppm | 4.3446 | 0.28862 | 3.7673 | 4.9218 | 3.4787 | 5.2104 | 6.643% | 13.29% | 19.93% | 4.1273 | 4.5618
Ho, ppm | 0.62908 | 0.034688 | 0.5597 | 0.69845 | 0.52501 | 0.73314 | 5.514% | 11.03% | 16.54% | 0.59762 | 0.66053
In, ppm | 0.058553 | 0.0044196 | 0.049714 | 0.067392 | 0.045295 | 0.071812 | 7.548% | 15.1% | 22.64% | 0.055626 | 0.061481
K, wt.% | 1.3185 | 0.051875 | 1.2148 | 1.4223 | 1.1629 | 1.4741 | 3.934% | 7.869% | 11.8% | 1.2526 | 1.3844
La, ppm | 37.9907 | 1.7448 | 34.5012 | 41.4803 | 32.7564 | 43.225 | 4.593% | 9.185% | 13.78% | 36.0912 | 39.8902
Li, ppm | 21.7332 | 1.0741 | 19.5851 | 23.8814 | 18.511 | 24.9554 | 4.942% | 9.884% | 14.83% | 20.6466 | 22.8199
Lu, ppm | 0.17769 | 0.022856 | 0.13198 | 0.2234 | 0.10913 | 0.24626 | 12.86% | 25.72% | 38.59% | 0.16881 | 0.18658
Mg, wt.% | 1.5542 | 0.049715 | 1.4548 | 1.6536 | 1.405 | 1.7033 | 3.199% | 6.398% | 9.596% | 1.4765 | 1.6319
Mn, wt.% | 0.049236 | 0.001489 | 0.046258 | 0.052214 | 0.044769 | 0.053703 | 3.024% | 6.048% | 9.072% | 0.046774 | 0.051698
Mo, ppm | 1.9642 | 0.16693 | 1.6304 | 2.2981 | 1.4635 | 2.465 | 8.498% | 17% | 25.49% | 1.866 | 2.0624
Na, wt.% | 1.0741 | 0.023425 | 1.0272 | 1.1209 | 1.0038 | 1.1443 | 2.181% | 4.362% | 6.543% | 1.0204 | 1.1278
Nb, ppm | 31.7716 | 1.7708 | 28.23 | 35.3132 | 26.4592 | 37.084 | 5.574% | 11.15% | 16.72% | 30.183 | 33.3602
Nd, ppm | 32.0964 | 1.7015 | 28.6934 | 35.4995 | 26.9919 | 37.201 | 5.301% | 10.6% | 15.9% | 30.4916 | 33.7013
Ni, ppm | 83.1882 | 4.0172 | 75.1538 | 91.2227 | 71.1365 | 95.24 | 4.829% | 9.658% | 14.49% | 79.0288 | 87.3477
P, wt.% | 0.095295 | 0.0034028 | 0.088489 | 0.1021 | 0.085086 | 0.1055 | 3.571% | 7.142% | 10.71% | 0.09053 | 0.10006
Pb, ppm | 9.9208 | 0.9145 | 8.0918 | 11.7498 | 7.1773 | 12.6643 | 9.218% | 18.44% | 27.65% | 9.4247 | 10.4168
Pr, ppm | 8.5113 | 0.46468 | 7.5819 | 9.4407 | 7.1173 | 9.9053 | 5.46% | 10.92% | 16.38% | 8.0857 | 8.9369
Rb, ppm | 64.7019 | 3.7977 | 57.1065 | 72.2972 | 53.3089 | 76.0949 | 5.869% | 11.74% | 17.61% | 61.4668 | 67.937
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.027184 | 0.003726 | 0.019732 | 0.034636 | 0.016006 | 0.038362 | 13.71% | 27.41% | 41.12% | 0.025825 | 0.028543
Sb, ppm | 0.88377 | 0.06807 | 0.74763 | 1.0199 | 0.67956 | 1.088 | 7.702% | 15.4% | 23.11% | 0.83958 | 0.92796
Sc, ppm | 9.8918 | 0.6239 | 8.644 | 11.1396 | 8.0201 | 11.7635 | 6.307% | 12.61% | 18.92% | 9.3972 | 10.3864
Sm, ppm | 6.2719 | 0.33604 | 5.5998 | 6.944 | 5.2638 | 7.28 | 5.358% | 10.72% | 16.07% | 5.9583 | 6.5855
Sn, ppm | 3.5909 | 0.16568 | 3.2595 | 3.9222 | 3.0939 | 4.0879 | 4.614% | 9.228% | 13.84% | 3.4113 | 3.7704
Sr, ppm | 311.4662 | 7.7313 | 296.0037 | 326.9288 | 288.2724 | 334.66 | 2.482% | 4.964% | 7.447% | 295.8929 | 327.0395
Ta, ppm | 2.0748 | 0.17486 | 1.7251 | 2.4245 | 1.5503 | 2.5994 | 8.428% | 16.86% | 25.28% | 1.9711 | 2.1786
Tb, ppm | 0.68798 | 0.01816 | 0.65166 | 0.7243 | 0.6335 | 0.74246 | 2.64% | 5.279% | 7.919% | 0.65358 | 0.72238
Th, ppm | 8.6755 | 0.34352 | 7.9884 | 9.3625 | 7.6449 | 9.7061 | 3.96% | 7.919% | 11.88% | 8.2417 | 9.1093
Ti, wt.% | 0.56692 | 0.014012 | 0.5389 | 0.59495 | 0.52489 | 0.60896 | 2.472% | 4.943% | 7.415% | 0.53858 | 0.59527
Tl, ppm | 0.35326 | 0.017664 | 0.31794 | 0.38859 | 0.30027 | 0.40626 | 5% | 10% | 15% | 0.3356 | 0.37093
Tm, ppm | 0.20127 | 0.012595 | 0.17609 | 0.22646 | 0.16349 | 0.23906 | 6.257% | 12.51% | 18.77% | 0.19121 | 0.21134
U, ppm | 1.6213 | 0.10684 | 1.4076 | 1.835 | 1.3008 | 1.9419 | 6.59% | 13.18% | 19.77% | 1.5403 | 1.7024
V, ppm | 69.932 | 2.5804 | 64.7712 | 75.0928 | 62.1908 | 77.6732 | 3.69% | 7.38% | 11.07% | 66.4354 | 73.4286
W, ppm | 2.0536 | 0.13673 | 1.7801 | 2.327 | 1.6434 | 2.4637 | 6.658% | 13.32% | 19.97% | 1.9509 | 2.1562
Y, ppm | 15.6643 | 0.96223 | 13.7398 | 17.5888 | 12.7776 | 18.551 | 6.143% | 12.29% | 18.43% | 14.8811 | 16.4475
Yb, ppm | 1.2817 | 0.062749 | 1.1562 | 1.4072 | 1.0935 | 1.47 | 4.896% | 9.791% | 14.69% | 1.2176 | 1.3458
Zn, ppm | 74.9419 | 3.3113 | 68.3194 | 81.5645 | 65.0081 | 84.8758 | 4.418% | 8.837% | 13.26% | 71.1948 | 78.689
Zr, ppm | 187.837 | 10.2739 | 167.2893 | 208.3847 | 157.0154 | 218.6586 | 5.47% | 10.94% | 16.41% | 178.4451 | 197.2288
Aqua Regia Digestion
Ag, ppm | 0.078574 | 0.007642 | 0.06329 | 0.093858 | 0.055648 | 0.1015 | 9.726% | 19.45% | 29.18% | 0.074645 | 0.082503
Al, wt.% | 1.6421 | 0.12176 | 1.3986 | 1.8856 | 1.2768 | 2.0074 | 7.415% | 14.83% | 22.24% | 1.56 | 1.7242
As, ppm | 43.4324 | 1.3337 | 40.765 | 46.0999 | 39.4312 | 47.4336 | 3.071% | 6.142% | 9.212% | 41.2608 | 45.604
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 94.5276 | 4.966 | 84.5956 | 104.4596 | 79.6295 | 109.4257 | 5.254% | 10.51% | 15.76% | 89.8012 | 99.254
Be, ppm | 1.0776 | 0.06931 | 0.93898 | 1.2162 | 0.86967 | 1.2855 | 6.432% | 12.86% | 19.3% | 1.0237 | 1.1315
Bi, ppm | 0.2647 | 0.020776 | 0.22315 | 0.30625 | 0.20237 | 0.32703 | 7.849% | 15.7% | 23.55% | 0.25146 | 0.27793
Ca, wt.% | 0.52434 | 0.04097 | 0.4424 | 0.60628 | 0.40143 | 0.64725 | 7.814% | 15.63% | 23.44% | 0.49812 | 0.55056
Cd, ppm | 0.06726 | 0.023337 | 0.020585 | 0.11393 | 0 | 0.13727 | 34.7% | 69.39% | 104.1% | 0.063897 | 0.070623
Ce, ppm | 40.8783 | 1.8744 | 37.1294 | 44.6271 | 35.255 | 46.5016 | 4.585% | 9.171% | 13.76% | 38.8343 | 42.9222
Co, ppm | 15.9278 | 0.71454 | 14.4987 | 17.3568 | 13.7842 | 18.0714 | 4.486% | 8.972% | 13.46% | 15.1314 | 16.7242
Cr, ppm | 45.8997 | 5.4293 | 35.041 | 56.7583 | 29.6117 | 62.1876 | 11.83% | 23.66% | 35.49% | 43.6047 | 48.1947
Cs, ppm | 1.1273 | 0.092659 | 0.94196 | 1.3126 | 0.8493 | 1.4053 | 8.22% | 16.44% | 24.66% | 1.0709 | 1.1836
Cu, ppm | 18.8888 | 2.1742 | 14.5405 | 23.2372 | 12.3664 | 25.4113 | 11.51% | 23.02% | 34.53% | 17.9444 | 19.8333
Fe, wt.% | 3.0079 | 0.26284 | 2.4822 | 3.5335 | 2.2194 | 3.7964 | 8.738% | 17.48% | 26.21% | 2.8575 | 3.1583
Ga, ppm | 5.9673 | 0.52072 | 4.9259 | 7.0088 | 4.4052 | 7.5295 | 8.726% | 17.45% | 26.18% | 5.669 | 6.2657
Ge, ppm | 0.099667 | 0.016914 | 0.065838 | 0.1335 | 0.048924 | 0.15041 | 16.97% | 33.94% | 50.91% | 0.094683 | 0.10465
In, ppm | 0.025375 | 0.0034147 | 0.018546 | 0.032204 | 0.015131 | 0.035619 | 13.46% | 26.91% | 40.37% | 0.024106 | 0.026644
K, wt.% | 0.33493 | 0.024826 | 0.28528 | 0.38459 | 0.26046 | 0.40941 | 7.412% | 14.82% | 22.24% | 0.31819 | 0.35168
La, ppm | 23.162 | 0.96539 | 21.2312 | 25.0928 | 20.2658 | 26.0582 | 4.168% | 8.336% | 12.5% | 22.0039 | 24.3201
Li, ppm | 7.6061 | 0.62756 | 6.351 | 8.8612 | 5.7234 | 9.4888 | 8.251% | 16.5% | 24.75% | 7.2258 | 7.9864
Mg, wt.% | 0.98857 | 0.068719 | 0.85114 | 1.126 | 0.78242 | 1.1947 | 6.951% | 13.9% | 20.85% | 0.93915 | 1.038
Mn, wt.% | 0.033367 | 0.0023371 | 0.028693 | 0.038041 | 0.026356 | 0.040378 | 7.004% | 14.01% | 21.01% | 0.031699 | 0.035036
Mo, ppm | 1.3238 | 0.11847 | 1.0868 | 1.5607 | 0.96836 | 1.6792 | 8.95% | 17.9% | 26.85% | 1.2576 | 1.39
Na, wt.% | 0.35057 | 0.021956 | 0.30666 | 0.39448 | 0.2847 | 0.41644 | 6.263% | 12.53% | 18.79% | 0.33304 | 0.3681
Nb, ppm | 0.58659 | 0.094041 | 0.39851 | 0.77467 | 0.30446 | 0.86871 | 16.03% | 32.06% | 48.1% | 0.55726 | 0.61592
Ni, ppm | 72.6347 | 3.4325 | 65.7697 | 79.4996 | 62.3373 | 82.932 | 4.726% | 9.451% | 14.18% | 69.0029 | 76.2664
P, wt.% | 0.05846 | 0.0038009 | 0.050859 | 0.066062 | 0.047058 | 0.069863 | 6.502% | 13% | 19.5% | 0.055537 | 0.061383
Pb, ppm | 7.0819 | 0.55574 | 5.9704 | 8.1934 | 5.4147 | 8.7491 | 7.847% | 15.69% | 23.54% | 6.7278 | 7.436
Rb, ppm | 19.0763 | 1.1575 | 16.7614 | 21.3912 | 15.6039 | 22.5487 | 6.068% | 12.14% | 18.2% | 18.1225 | 20.0301
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.02435 | 0.0042528 | 0.015845 | 0.032856 | 0.011592 | 0.037109 | 17.46% | 34.93% | 52.39% | 0.023133 | 0.025568
Sb, ppm | 0.44672 | 0.058316 | 0.33009 | 0.56335 | 0.27177 | 0.62167 | 13.05% | 26.11% | 39.16% | 0.42439 | 0.46906
Sc, ppm | 3.7374 | 0.31432 | 3.1087 | 4.366 | 2.7944 | 4.6803 | 8.41% | 16.82% | 25.23% | 3.5505 | 3.9242
Sn, ppm | 1.2367 | 0.13078 | 0.9751 | 1.4982 | 0.84432 | 1.629 | 10.58% | 21.15% | 31.73% | 1.1748 | 1.2985
Sr, ppm | 63.4029 | 2.589 | 58.225 | 68.5809 | 55.6361 | 71.1698 | 4.083% | 8.167% | 12.25% | 60.2328 | 66.5731
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.44776 | 0.034133 | 0.3795 | 0.51603 | 0.34536 | 0.55016 | 7.623% | 15.25% | 22.87% | 0.42537 | 0.47015
Th, ppm | 5.8503 | 0.24652 | 5.3573 | 6.3434 | 5.1108 | 6.5899 | 4.214% | 8.427% | 12.64% | 5.5578 | 6.1429
Ti, wt.% | 0.1622 | 0.012776 | 0.13665 | 0.18775 | 0.12387 | 0.20053 | 7.877% | 15.75% | 23.63% | 0.15409 | 0.17031
Tl, ppm | 0.1356 | 0.01583 | 0.10394 | 0.16726 | 0.088115 | 0.18309 | 11.67% | 23.35% | 35.02% | 0.12882 | 0.14238
U, ppm | 0.96523 | 0.078699 | 0.80783 | 1.1226 | 0.72913 | 1.2013 | 8.153% | 16.31% | 24.46% | 0.91697 | 1.0135
V, ppm | 33.6989 | 1.83 | 30.0388 | 37.359 | 28.2088 | 39.189 | 5.431% | 10.86% | 16.29% | 32.014 | 35.3838
W, ppm | 0.38335 | 0.038302 | 0.30675 | 0.45996 | 0.26844 | 0.49826 | 9.991% | 19.98% | 29.97% | 0.36418 | 0.40252
Y, ppm | 9.1007 | 0.38568 | 8.3293 | 9.872 | 7.9436 | 10.2577 | 4.238% | 8.476% | 12.71% | 8.6456 | 9.5557
Yb, ppm | 0.63153 | 0.040344 | 0.55085 | 0.71222 | 0.5105 | 0.75257 | 6.388% | 12.78% | 19.16% | 0.59996 | 0.66311
Zn, ppm | 46.3867 | 4.6986 | 36.9895 | 55.7839 | 32.2909 | 60.4825 | 10.13% | 20.26% | 30.39% | 44.0674 | 48.706
Zr, ppm | 37.1153 | 8.6424 | 19.8305 | 54.4001 | 11.1881 | 63.0425 | 23.29% | 46.57% | 69.86% | 35.2595 | 38.9711
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 250c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 250c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 250c (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 16 | 8 | 245400 | 0.08689 | 0.31453 | 0.32861
2 | 1 | ANSLu | 1 | 17 | 9 | 245401 | 0.087447 | 0.31602 | 0.32869
2 | 1 | ANSLu | 1 | 19 | 10 | 245402 | 0.087702 | 0.33842 | 0.32989
2 | 1 | ANSLu | 1 | 4 | 2 | 245403 | 0.086053 | 0.35105 | 0.33056
2 | 1 | ANSLu | 1 | 3 | 2 | 245404 | 0.084581 | 0.31832 | 0.32881
2 | 1 | ANSLu | 1 | 9 | 5 | 245405 | 0.085104 | 0.31026 | 0.32838
2 | 1 | ANSLu | 1 | 5 | 3 | 245406 | 0.087549 | 0.36617 | 0.33137
2 | 1 | ANSLu | 1 | 10 | 5 | 245407 | 0.087321 | 0.324 | 0.32912
2 | 1 | ANSLu | 1 | 20 | 10 | 245408 | 0.08611 | 0.32684 | 0.32927
2 | 1 | ANSLu | 1 | 7 | 4 | 245409 | 0.084814 | 0.32994 | 0.32943
2 | 1 | ANSLu | 1 | 18 | 9 | 245410 | 0.086898 | 0.31669 | 0.32872
2 | 1 | ANSLu | 1 | 13 | 7 | 245411 | 0.086577 | 0.32967 | 0.32942
2 | 1 | ANSLu | 1 | 15 | 8 | 245412 | 0.08398 | 0.32203 | 0.32901
2 | 1 | ANSLu | 1 | 12 | 6 | 245413 | 0.086145 | 0.33013 | 0.32944
2 | 1 | ANSLu | 1 | 14 | 7 | 245414 | 0.087215 | 0.33018 | 0.32945
2 | 1 | ANSLu | 1 | 1 | 1 | 245415 | 0.08658 | 0.3235 | 0.32909
2 | 1 | ANSLu | 1 | 11 | 6 | 245416 | 0.083989 | 0.33293 | 0.32959
2 | 1 | ANSLu | 1 | 8 | 4 | 245417 | 0.084456 | 0.31705 | 0.32874
2 | 1 | ANSLu | 1 | 2 | 1 | 245418 | 0.087195 | 0.33662 | 0.32979
2 | 1 | ANSLu | 1 | 6 | 3 | 245419 | 0.083587 | 0.35372 | 0.33071
Mean | [6 empty] | 0.08601 | 0.3294 | 0.3294
Median | [7 empty] | 0.32826 | 0.32934
SD | [7 empty] | 0.014303 | 0.00076587
RSD | [7 empty] | 4.342% | 0.2325%
RSD Upscale Factor | [7 empty] | 0.053544
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 250c (Certified Value 0.313 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | FA*MS | FA*MS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*AAS | 3
|  |  | 0.085g | 25g | 50g | 30g | 50g | 30g | 50g | 50g | 50g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 0.3235 | 0.321 | 0.30952 | 0.31867 | 0.32 | 0.316 | 0.3 | 0.326 | 0.274 | 0.302 | 0.3 | 0.344 | 0.305 | 0.319 | 0.32 | 0.31 | 0.303 | 0.305 | 0.311 | 0.306 | 0.307 | 0.31 | 1
| 1 | 2 | 0.33662 | 0.323 | 0.31894 | 0.32333 | 0.33 | 0.337 | 0.31 | 0.323 | 0.294 | 0.302 | 0.3 | 0.34247 | 0.305 | 0.324 | 0.319 | 0.318 | 0.30262 | 0.312 | 0.312 | 0.299 | 0.309 | 0.3 | #N/A
| 1 | 3 | 0.31832 | 0.322 | 0.31517 | 0.31558 | 0.38 | 0.325 | 0.32 | 0.33 | 0.2755 | 0.265 | 0.3 | 0.349 | 0.303 | 0.326 | 0.322 | 0.304 | 0.3095 | 0.308 | 0.316 | 0.312 | 0.31 | 0.31 | 16
| 1 | 4 | 0.35105 | 0.324 | 0.31328 | 0.31749 | 0.36 | 0.34 | 0.27 | 0.321 | 0.287 | 0.299 | 0.3 | 0.343 | 0.306 | 0.328 | 0.321 | 0.315 | 0.30694 | 0.303 | 0.311 | 0.307 | 0.31 | 0.31 | 0.3125
| 1 | 5 | 0.36617 | 0.319 | 0.3114 | 0.31666 | 0.3 | 0.316 | 0.29 | 0.324 | 0.276 | 0.279 | 0.31 | 0.34 | 0.303 | 0.319 | 0.321 | 0.315 | 0.31313 | 0.314 | 0.316 | 0.306 | 0.309 | 0.31 | 7
| 1 | 6 | 0.35372 | 0.327 | 0.31988 | 0.31665 | 0.34 | 0.303 | 0.25 | 0.324 | 0.279 | 0.31 | 0.3 | 0.34083 | 0.323 | 0.327 | 0.318 | 0.315 | 0.30956 | 0.308 | 0.302 | 0.296 | 0.308 | 0.31
|  | 7 | 0.32994
|  | 8 | 0.31705
|  | 9 | 0.31026
|  | 10 | 0.324
|  | 11 | 0.33293
|  | 12 | 0.33013
|  | 13 | 0.32967
|  | 14 | 0.33018
|  | 15 | 0.32203
|  | 16 | 0.31453
|  | 17 | 0.31602
|  | 18 | 0.31669
|  | 19 | 0.33842
|  | 20 | 0.32684
| Mean |  | 0.3294 | 0.32267 | 0.3147 | 0.31806 | 0.33833 | 0.32283 | 0.29 | 0.32467 | 0.28092 | 0.29283 | 0.30167 | 0.34322 | 0.3075 | 0.32383 | 0.32017 | 0.31283 | 0.30746 | 0.30833 | 0.31133 | 0.30433 | 0.30883 | 0.30833
| Median |  | 0.32826 | 0.3225 | 0.31423 | 0.31708 | 0.335 | 0.3205 | 0.295 | 0.324 | 0.2775 | 0.3005 | 0.3 | 0.34274 | 0.305 | 0.325 | 0.3205 | 0.315 | 0.30822 | 0.308 | 0.3115 | 0.306 | 0.309 | 0.31
| Std Dev. |  | 0.014303 | 0.0027325 | 0.0041162 | 0.0027751 | 0.028577 | 0.014049 | 0.026077 | 0.0030768 | 0.0079146 | 0.017128 | 0.0040825 | 0.003184 | 0.0076877 | 0.0039707 | 0.001472 | 0.0050365 | 0.004105 | 0.0041312 | 0.0051251 | 0.0058195 | 0.001169 | 0.0040825
| Rel.Std.Dev. |  | 4.342% | 0.8469% | 1.308% | 0.8725% | 8.447% | 4.352% | 8.992% | 0.9477% | 2.817% | 5.849% | 1.353% | 0.9277% | 2.5% | 1.226% | 0.4597% | 1.61% | 1.335% | 1.34% | 1.646% | 1.912% | 0.3785% | 1.324%
| PDM3 |  | 5.409% | 3.253% | 0.7029% | 1.78% | 8.266% | 3.306% | -7.2% | 3.893% | -10.11% | -6.294% | -3.467% | 9.829% | -1.6% | 3.626% | 2.453% | 0.1064% | -1.614% | -1.334% | -0.3736% | -2.614% | -1.174% | -1.334%
| Z-Score (Absolute) |  | NA | 0.87 | 0.26 | 0.51 | 2.06 | 0.88 | 1.63 | 1.02 | 2.32 | 1.41 | 0.74 | 2.43 | 0.29 | 0.95 | 0.67 | 0.11 | 0.3 | 0.23 | 0 | 0.53 | 0.19 | 0.23
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 250c (Certified Value 0.313 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | 3
|  |  | 0.085g | 25g | 30g | 50g | 30g | 50g | 30g | 30g | 25g | 15g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 3
| 1 | 1 | 0.3235 | 0.32 | 0.33281 | 0.3 | 0.32 | 0.29333 | 0.289 | 0.316 | 0.308 | 0.3045 | 0.307 | 0.291 | 0.2984 | 0.322 | 0.3 | 0.306 | 0.34 | 1
| 1 | 2 | 0.33662 | 0.3 | 0.34294 | 0.3 | 0.32 | 0.28667 | 0.334 | 0.319 | 0.306 | 0.3045 | 0.301 | 0.305 | 0.29691 | 0.328 | 0.305 | 0.308 | 0.33 | #N/A
| 1 | 3 | 0.31832 | 0.31 | 0.33889 | 0.36 | 0.32 | 0.28 | 0.284 | 0.337 | 0.314 | 0.3045 | 0.307 | 0.299 | 0.30188 | 0.318 | 0.296 | 0.309 | 0.36 | 16
| 1 | 4 | 0.35105 | 0.32 | 0.33687 | 0.34 | 0.32 | 0.27333 | 0.284 | 0.315 | 0.316 | 0.3255 | 0.3 | 0.31 | 0.30327 | 0.323 | 0.305 | 0.311 | 0.34 | 0.31298
| 1 | 5 | 0.36617 | 0.32 | 0.33484 | 0.28 | 0.32 | 0.28 | 0.303 | 0.333 | 0.316 | 0.3045 | 0.3 | 0.3 | 0.30534 | 0.32 | 0.311 | 0.314 | 0.34 | 9
| 1 | 6 | 0.35372 | 0.31 | 0.34395 | 0.32 | 0.32 | 0.27333 | 0.324 | 0.334 | 0.32 | 0.336 | 0.304 | 0.299 | 0.29754 | 0.327 | 0.302 | 0.306 | 0.35
|  | 7 | 0.32994
|  | 8 | 0.31705
|  | 9 | 0.31026
|  | 10 | 0.324
|  | 11 | 0.33293
|  | 12 | 0.33013
|  | 13 | 0.32967
|  | 14 | 0.33018
|  | 15 | 0.32203
|  | 16 | 0.31453
|  | 17 | 0.31602
|  | 18 | 0.31669
|  | 19 | 0.33842
|  | 20 | 0.32684
| Mean |  | 0.3294 | 0.31333 | 0.33838 | 0.31667 | 0.32 | 0.28111 | 0.303 | 0.32567 | 0.31333 | 0.31325 | 0.30317 | 0.30067 | 0.30055 | 0.323 | 0.30317 | 0.309 | 0.34333
| Median |  | 0.32826 | 0.315 | 0.33788 | 0.31 | 0.32 | 0.28 | 0.296 | 0.326 | 0.315 | 0.3045 | 0.3025 | 0.2995 | 0.30014 | 0.3225 | 0.3035 | 0.3085 | 0.34
| Std Dev. |  | 0.014303 | 0.008165 | 0.0044261 | 0.029439 | 0 | 0.0077936 | 0.021541 | 0.010033 | 0.0053166 | 0.013956 | 0.0033116 | 0.0064083 | 0.0034369 | 0.0038987 | 0.0051153 | 0.0030984 | 0.010328
| Rel.Std.Dev. |  | 4.342% | 2.606% | 1.308% | 9.297% | 0% | 2.772% | 7.109% | 3.081% | 1.697% | 4.455% | 1.092% | 2.131% | 1.144% | 1.207% | 1.687% | 1.003% | 3.008%
| PDM3 |  | 5.249% | 0.1138% | 8.118% | 1.179% | 2.244% | -10.18% | -3.188% | 4.054% | 0.1138% | 0.0872% | -3.135% | -3.933% | -3.969% | 3.202% | -3.135% | -1.271% | 9.699%
| Z-Score (Absolute) |  | NA | 0 | 1.67 | 0.22 | 0.45 | 2.14 | 0.69 | 0.82 | 0 | 0 | 0.67 | 0.84 | 0.85 | 0.65 | 0.67 | 0.29 | 2
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 250c (Certified Value 0.295 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*OES | CNL*MS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 05g | 50g | 50g | 200g | 50g | 30g | 30g | 30g | 25g | 20g | 05g | 3
| 1 | 1 | 0.3235 | 0.31 | 0.29504 | 0.26127 | 0.28 | 0.3 | 0.3 | 0.26942 | 0.283 | 0.29 | 0.303 | 0.2499 | 0.3 | 0.298 | 0.3176 | 0.3037 | 0.307 | 1
| 1 | 2 | 0.33662 | 0.32 | 0.30402 | 0.26192 | 0.27 | 0.28 | 0.3 | 0.29186 | 0.286 | 0.3 | 0.312 | 0.2598 | 0.31 | 0.302 | 0.3097 | 0.29359 | 0.296 | #N/A
| 1 | 3 | 0.31832 | 0.32 | 0.30043 | 0.2704 | 0.31 | 0.36 | 0.3 | 0.27437 | 0.298 | 0.3 | 0.312 | 0.2519 | 0.31 | 0.301 | 0.3128 | 0.287 | 0.312 | 16
| 1 | 4 | 0.35105 | 0.31 | 0.29863 | 0.26423 | 0.32 | 0.28 | 0.3 | 0.26642 | 0.304 | 0.29 | 0.299 | 0.2514 | 0.3 | 0.301 | 0.3188 | 0.28909 | 0.311 | 0.29515
| 1 | 5 | 0.36617 | 0.31 | 0.29684 | 0.27025 | 0.27 | 0.29 | 0.3 | 0.2693 | 0.284 | 0.3 | 0.3 | 0.2541 | 0.3 | 0.3 | 0.3021 | 0.30915 | 0.307 | 11
| 1 | 6 | 0.35372 | 0.31 | 0.30492 | 0.26392 | 0.25 | 0.28 | 0.3 | 0.26656 | 0.286 | 0.3 | 0.299 | 0.251 | 0.31 | 0.299 | 0.31 | 0.2977 | 0.306
|  | 7 | 0.32994
|  | 8 | 0.31705
|  | 9 | 0.31026
|  | 10 | 0.324
|  | 11 | 0.33293
|  | 12 | 0.33013
|  | 13 | 0.32967
|  | 14 | 0.33018
|  | 15 | 0.32203
|  | 16 | 0.31453
|  | 17 | 0.31602
|  | 18 | 0.31669
|  | 19 | 0.33842
|  | 20 | 0.32684
| Mean |  | 0.3294 | 0.31333 | 0.29998 | 0.26533 | 0.28333 | 0.29833 | 0.3 | 0.27299 | 0.29017 | 0.29667 | 0.30417 | 0.25302 | 0.305 | 0.30017 | 0.31183 | 0.2967 | 0.3065
| Median |  | 0.32826 | 0.31 | 0.29953 | 0.26408 | 0.275 | 0.285 | 0.3 | 0.26936 | 0.286 | 0.3 | 0.3015 | 0.25165 | 0.305 | 0.3005 | 0.3114 | 0.29564 | 0.307
| Std Dev. |  | 0.014303 | 0.005164 | 0.0039237 | 0.0040312 | 0.026583 | 0.031252 | 0 | 0.009681 | 0.0086814 | 0.005164 | 0.0062423 | 0.0036008 | 0.0054772 | 0.001472 | 0.0060889 | 0.0085655 | 0.0056833
| Rel.Std.Dev. |  | 4.342% | 1.648% | 1.308% | 1.519% | 9.382% | 10.48% | 0% | 3.546% | 2.992% | 1.741% | 2.052% | 1.423% | 1.796% | 0.4904% | 1.953% | 2.887% | 1.854%
| PDM3 |  | 11.6% | 6.159% | 1.634% | -10.1% | -4.005% | 1.077% | 1.642% | -7.51% | -1.69% | 0.5122% | 3.053% | -14.28% | 3.336% | 1.698% | 5.651% | 0.5249% | 3.844%
| Z-Score (Absolute) |  | NA | 1.24 | 0.14 | 2.72 | 1.24 | 0 | NA | 2.09 | 0.67 | 0.14 | 0.48 | 3.74 | 0.55 | 0.15 | 1.11 | 0.13 | 0.67
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 250c (Certified Value 0.311 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | [11 empty] | 3
| 1 | 1 | 0.3235 | 0.321 | 0.293 | 0.32 | 0.27 | 0.29 | 0.28 | 0.31 | 0.313 | 0.293 | 0.268 | 0.355 | 1
| 1 | 2 | 0.33662 | 0.312 | 0.277 | 0.33 | 0.31 | 0.3 | 0.32 | 0.28 | 0.307 | 0.313 | 0.325 | 0.292 | #N/A
| 1 | 3 | 0.31832 | 0.345 | 0.287 | 0.35 | 0.32 | 0.35 | 0.31 | 0.32 | 0.354 | 0.324 | 0.279 | 0.366 | 16
| 1 | 4 | 0.35105 | 0.353 | 0.332 | 0.33 | 0.31 | 0.36 | 0.29 | 0.28 | 0.284 | 0.33 | 0.331 | 0.276 | 0.31092
| 1 | 5 | 0.36617 | 0.311 | 0.292 | 0.31 | 0.3 | 0.27 | 0.31 | 0.31 | 0.298 | 0.272 | 0.293 | 0.388 | 13
| 1 | 6 | 0.35372 | 0.339 | 0.281 | 0.3 | 0.3 | 0.33 | 0.31 | 0.31 | 0.365 | 0.315 | 0.274 | 0.283
|  | 7 | 0.32994
|  | 8 | 0.31705
|  | 9 | 0.31026
|  | 10 | 0.324
|  | 11 | 0.33293
|  | 12 | 0.33013
|  | 13 | 0.32967
|  | 14 | 0.33018
|  | 15 | 0.32203
|  | 16 | 0.31453
|  | 17 | 0.31602
|  | 18 | 0.31669
|  | 19 | 0.33842
|  | 20 | 0.32684
| Mean |  | 0.3294 | 0.33017 | 0.29367 | 0.32333 | 0.30167 | 0.31667 | 0.30333 | 0.30167 | 0.32017 | 0.30783 | 0.295 | 0.32667
| Median |  | 0.32826 | 0.33 | 0.2895 | 0.325 | 0.305 | 0.315 | 0.31 | 0.31 | 0.31 | 0.314 | 0.286 | 0.3235
| Std Dev. |  | 0.014303 | 0.017893 | 0.019775 | 0.017512 | 0.017224 | 0.03559 | 0.015055 | 0.017224 | 0.032183 | 0.021609 | 0.02693 | 0.048554
| Rel.Std.Dev. |  | 4.342% | 5.419% | 6.734% | 5.416% | 5.71% | 11.24% | 4.963% | 5.71% | 10.05% | 7.02% | 9.129% | 14.86%
| PDM3 |  | 5.944% | 6.189% | -5.55% | 3.991% | -2.977% | 1.847% | -2.441% | -2.977% | 2.973% | -0.9941% | -5.122% | 5.063%
| Z-Score (Absolute) |  | NA | 1.22 | 0.77 | 0.85 | 0.34 | 0.48 | 0.25 | 0.34 | 0.67 | 0 | 0.7 | 1.03
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 250c (Certified Value < 0.1 ppm) | [19 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.1 | < 0.2 | < 0.1 | < 0.1 | < 0.3 | 0.18 | < 0.5 | 0.14 | < 0.1 | < 0.3 | 0.1 | < 0.1 | 0.3 | 0.08 | 0.07 | 0.06 |  |  | 1
| 1 | 2 | 0.09 | < 0.2 | < 0.1 | < 0.1 | < 0.3 | 0.18 | < 0.5 | 0.09 | < 0.1 | < 0.3 | 0.1 | < 0.1 | 0.3 | 0.05 | 0.04 | 0.06 |  |  | 20
| 1 | 3 | 0.1 | < 0.2 | < 0.1 | < 0.1 | < 0.3 | 0.18 | < 0.5 | 0.14 | < 0.1 | < 0.3 | 0.09 | < 0.1 | 0.3 | 0.06 | 0.06 | 0.09 |  |  | 16
| 1 | 4 | 0.09 | < 0.2 | < 0.1 | < 0.1 | < 0.3 | 0.19 | < 0.5 | 0.13 | < 0.1 | < 0.3 | 0.07 | < 0.1 | 0.3 | 0.06 | 0.05 | 0.08 |  |  | < 0.1
| 1 | 5 | 0.09 | < 0.2 | < 0.1 | < 0.1 | < 0.3 | 0.19 | < 0.5 | 0.1 | < 0.1 | < 0.3 | 0.09 | < 0.1 | 0.3 | 0.08 | 0.05 | 0.09 |  |  | 15
| 1 | 6 | 0.08 | < 0.2 | < 0.1 | < 0.1 | < 0.3 | 0.18 | < 0.5 | 0.11 | < 0.1 | < 0.3 | 0.11 | < 0.1 | 0.3 | 0.07 | 0.05 | 0.08
| Mean |  | 0.091667 |  |  |  |  | 0.18333 |  | 0.11833 |  |  | 0.093333 |  | 0.3 | 0.066667 | 0.053333 | 0.076667
| Median |  | 0.09 |  |  |  |  | 0.18 |  | 0.12 |  |  | 0.095 |  | 0.3 | 0.065 | 0.05 | 0.08
| Std Dev. |  | 0.0075277 |  |  |  |  | 0.005164 |  | 0.02137 |  |  | 0.013663 |  | 0 | 0.012111 | 0.010328 | 0.013663
| Rel.Std.Dev. |  | 8.212% |  |  |  |  | 2.817% |  | 18.06% |  |  | 14.64% |  | 0% | 18.17% | 19.36% | 17.82%
PDM3
| Z-Score (Absolute) |  | 0.01 | 0.12 | 0.67 | 0.67 | 0.91 | 1.44 | 2.5 | 0.41 | 0.67 | 0.91 | 0.01 | 0.67 | 3.29 | 0.41 | 0.62 | 0.25
| Analytical results for Al in OREAS 250c (Certified Value 5.18 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.2828 | 5.2707 | 5.5672 | 4.9135 | 5.15 | 5.55 | 5.26 | 5.25 | 5.01 | 5.01 | 5.1063 | 5.3 | 5.167 | 5.38 | 5.06 | 5.03 | 5.16 | 5.16 | 1
| 1 | 2 | 5.1403 | 5.2685 | 5.5802 | 4.9224 | 5.19 | 5.64 | 5.33 | 5.17 | 4.99 | 5.14 | 4.9169 | 5.11 | 5.1977 | 5.41 | 5.06 | 5.12 | 5.18 | 5.14 | #N/A
| 1 | 3 | 5.1263 | 5.2367 | 5.6024 | 4.823 | 5.24 | 5.57 | 5.23 | 5.13 | 5.28 | 5.06 | 5.092 | 5.2 | 5.2344 | 5.36 | 5.09 | 4.98 | 5 | 5.03 | 16
| 1 | 4 | 5.2337 | 5.2054 | 5.4881 | 4.8356 | 5.19 | 5.58 | 5.24 | 5.21 | 5.17 | 5.12 | 5.1041 | 5.14 | 5.2266 | 5.55 | 5.2 | 5.09 | 5.05 | 5.06 | 5.176
| 1 | 5 | 5.3238 | 5.2707 | 5.5085 | 4.9254 | 5.18 | 5.53 | 5.28 | 5.16 | 5.09 | 5.11 | 5.0963 | 5.29 | 5.1805 | 5.29 | 5.2 | 5.06 | 5.32 | 5.13 | 16
| 1 | 6 | 4.9446 | 5.2174 | 5.575 | 4.9246 | 5.14 | 5.52 | 5.32 | 5.28 | 5.12 | 5.11 | 5.1059 | 5.31 | 5.199 | 5.41 | 5.24 | 5.18 | 5.17 | 5.07
| Mean |  | 5.1753 | 5.2449 | 5.5536 | 4.8908 | 5.1817 | 5.565 | 5.2767 | 5.2 | 5.11 | 5.0917 | 5.0702 | 5.225 | 5.2009 | 5.4 | 5.1417 | 5.0767 | 5.1467 | 5.0983
| Median |  | 5.187 | 5.2526 | 5.5711 | 4.9179 | 5.185 | 5.56 | 5.27 | 5.19 | 5.105 | 5.11 | 5.1002 | 5.245 | 5.1983 | 5.395 | 5.145 | 5.075 | 5.165 | 5.1
| Std Dev. |  | 0.13702 | 0.029219 | 0.044839 | 0.047967 | 0.035449 | 0.043243 | 0.041312 | 0.057271 | 0.10714 | 0.047924 | 0.075345 | 0.08735 | 0.025932 | 0.08579 | 0.080602 | 0.070048 | 0.11201 | 0.051929
| Rel.Std.Dev. |  | 2.648% | 0.5571% | 0.8074% | 0.9808% | 0.6841% | 0.7771% | 0.7829% | 1.101% | 2.097% | 0.9412% | 1.486% | 1.672% | 0.4986% | 1.589% | 1.568% | 1.38% | 2.176% | 1.019%
| PDM3 |  | -0.0148% | 1.331% | 7.294% | -5.511% | 0.1092% | 7.515% | 1.945% | 0.4634% | -1.275% | -1.63% | -2.043% | 0.9464% | 0.4799% | 4.327% | -0.6636% | -1.919% | -0.567% | -1.501%
| Z-Score (Absolute) |  | 0.03 | 0.6 | 3.4 | 2.61 | 0.03 | 3.51 | 0.89 | 0.2 | 0.62 | 0.79 | 0.98 | 0.42 | 0.2 | 2.01 | 0.33 | 0.92 | 0.29 | 0.73
| Analytical results for As in OREAS 250c (Certified Value 54 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
1
| 1 | 1 | 54.9 | 61.246 | 114.6881 | 52 | 41.3 | 52 | 52.4767 | 54.7 | 53 | 39.8714 | 53.2 | 52.3239 | 54.3 | 55.3 | 56 | 57.4 | 53 |  | 1
| 1 | 2 | 56.5 | 60.8978 | 91.4365 | 53 | 35.8 | 51 | 51.925 | 54.7 | 57 | 39.7829 | 51.9 | 50.0359 | 53.5 | 55.7 | 57.6 | 56.8 | 52 |  | 21
| 1 | 3 | 56.4 | 60.2355 | 136.5034 | 53 | 35.7 | 51 | 50.75 | 58.3 | 54 | 39.9635 | 52.6 | 51.0564 | 53.5 | 55.9 | 57.7 | 56.9 | 53 |  | 16
| 1 | 4 | 55.9 | 59.7317 | 141.1511 | 52 | 35.3 | 50 | 53.8533 | 56.4 | 54 | 39.8098 | 48.3 | 50.7879 | 53.7 | 56.9 | 56.5 | 57.2 | 54 |  | 54.0987
| 1 | 5 | 54.3 | 60.5858 | 136.2342 | 52 | 39.8 | 53 | 51.5433 | 55.7 | 53 | 39.8462 | 51.1 | 53.7893 | 52.3 | 57.3 | 54.5 | 58.5 | 52 |  | 17
| 1 | 6 | 53 | 60.5137 | 138.0695 | 53 | 39.7 | 53 | 52.2333 | 55.4 | 55 | 39.498 | 52.6 | 53.6451 | 53.8 | 58 | 57 | 59.1 | 55
| Mean |  | 55.1667 | 60.5351 | 126.3471 | 52.5 | 37.9333 | 51.6667 | 52.1303 | 55.8667 | 54.3333 | 39.7953 | 51.6167 | 51.9398 | 53.5167 | 56.5167 | 56.55 | 57.65 | 53.1667
| Median |  | 55.4 | 60.5498 | 136.3688 | 52.5 | 37.75 | 51.5 | 52.0792 | 55.55 | 54 | 39.828 | 52.25 | 51.6902 | 53.6 | 56.4 | 56.75 | 57.3 | 53
| Std Dev. |  | 1.3677 | 0.52444 | 19.5559 | 0.54772 | 2.6235 | 1.2111 | 1.0384 | 1.3545 | 1.5055 | 0.15838 | 1.777 | 1.5626 | 0.66458 | 1.0477 | 1.1946 | 0.93541 | 1.169
| Rel.Std.Dev. |  | 2.479% | 0.8663% | 15.48% | 1.043% | 6.916% | 2.344% | 1.992% | 2.425% | 2.771% | 0.398% | 3.443% | 3.008% | 1.242% | 1.854% | 2.112% | 1.623% | 2.199%
| PDM3 |  | 1.974% | 11.9% | 133.5% | -2.955% | -29.88% | -4.496% | -3.639% | 3.268% | 0.4337% | -26.44% | -4.588% | -3.991% | -1.076% | 4.47% | 4.531% | 6.564% | -1.723%
| Z-Score (Absolute) |  | 0.59 | 2.49 | 25.85 | 0.36 | 5.53 | 0.66 | 0.49 | 0.83 | 0.29 | 4.87 | 0.67 | 0.56 | 0 | 1.06 | 1.08 | 1.47 | 0.12
| Analytical results for Ba in OREAS 250c (Certified Value 354 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 356.5 | 349.569 | 354.6982 | 332.8523 | 355 | 334 | 349 | 357 | 350 | 225 | 355.3098 | 360 | 353.8381 | 356 | 350 | 360 | 350 | 347 | 1
| 1 | 2 | 361.8 | 363.66 | 350.0516 | 332.9297 | 354 | 358 | 352 | 346 | 340 | 231 | 347.4082 | 350 | 343.7744 | 354 | 350 | 370 | 350 | 348 | 22
| 1 | 3 | 364.6 | 359.238 | 357.4727 | 326.8746 | 362 | 360 | 351 | 347 | 360 | 225 | 358.4555 | 360 | 345.2752 | 354 | 350 | 350 | 340 | 354 | 16
| 1 | 4 | 363.5 | 360.734 | 353.8603 | 327.5765 | 351 | 330 | 353 | 358 | 350 | 231 | 360.3329 | 350 | 351.6101 | 359 | 360 | 360 | 350 | 340 | 354.248
| 1 | 5 | 357 | 361.237 | 349.423 | 331.8836 | 354 | 323 | 352 | 352 | 350 | 234 | 354.6879 | 360 | 360.6398 | 345 | 360 | 360 | 360 | 339 | 18
| 1 | 6 | 344.3 | 373.142 | 351.9404 | 334.1952 | 350 | 324 | 351 | 356 | 350 | 233 | 357.0716 | 360 | 360.194 | 358 | 370 | 370 | 350 | 347
| Mean |  | 357.95 | 361.2633 | 352.9077 | 331.052 | 354.3333 | 338.1667 | 351.3333 | 352.6667 | 350 | 229.8333 | 355.5443 | 356.6667 | 352.5553 | 354.3333 | 356.6667 | 361.6667 | 350 | 345.8333
| Median |  | 359.4 | 360.9855 | 352.9003 | 332.368 | 354 | 332 | 351.5 | 354 | 350 | 231 | 356.1907 | 360 | 352.7241 | 355 | 355 | 360 | 350 | 347
| Std Dev. |  | 7.4723 | 7.5893 | 3.0394 | 3.0616 | 4.2269 | 16.6423 | 1.3663 | 5.2026 | 6.3246 | 3.92 | 4.4883 | 5.164 | 7.1603 | 5.0067 | 8.165 | 7.5277 | 6.3246 | 5.5648
| Rel.Std.Dev. |  | 2.088% | 2.101% | 0.8612% | 0.9248% | 1.193% | 4.921% | 0.3889% | 1.475% | 1.807% | 1.706% | 1.262% | 1.448% | 2.031% | 1.413% | 2.289% | 2.081% | 1.807% | 1.609%
| PDM3 |  | 1.045% | 1.98% | -0.3784% | -6.548% | 0.02408% | -4.54% | -0.8228% | -0.4464% | -1.199% | -35.12% | 0.3659% | 0.6827% | -0.4779% | 0.02408% | 0.6827% | 2.094% | -1.199% | -2.375%
| Z-Score (Absolute) |  | 1.04 | 1.71 | 0.02 | 4.4 | 0.31 | 2.96 | 0.29 | 0.02 | 0.56 | 24.87 | 0.56 | 0.78 | 0.05 | 0.31 | 0.78 | 1.8 | 0.56 | 1.41
| Analytical results for Be in OREAS 250c (Certified Value 1.97 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.98 | 2.0408 | 1.4886 | 1.8 | 1.9 | 1.95 | < 5 | 1.92 | 2 | 1.6615 | 1.99 | 2.0579 | 2.1 | 1.92 | 1.91 | 1.85 | 2.42 |  | 1
| 1 | 2 | 2.05 | 2.1147 | 1.5226 | 1.7 | 2.2 | 1.94 | < 5 | 1.88 | 2 | 1.6369 | 1.92 | 2.0426 | 2 | 1.9 | 1.95 | 1.86 | 2.31 |  | 23
| 1 | 3 | 2.08 | 2.1111 | 1.4286 | 1.7 | 2.1 | 1.93 | < 5 | 1.97 | 2 | 1.6307 | 1.99 | 2.1102 | 2 | 1.94 | 1.89 | 1.81 | 2.38 |  | 16
| 1 | 4 | 2.03 | 2.1107 | 1.4349 | 1.6 | 2 | 1.93 | < 5 | 1.96 | 2 | 1.6839 | 1.95 | 1.8582 | 2.1 | 1.96 | 1.94 | 1.83 | 2.26 |  | 1.9732
| 1 | 5 | 1.92 | 2.0884 | 1.4997 | 1.6 | 2 | 1.99 | < 5 | 1.9 | 2 | 1.6856 | 2 | 2.1798 | 1.9 | 1.98 | 1.93 | 1.92 | 2.31 |  | 19
| 1 | 6 | 1.83 | 2.0613 | 1.5243 | 1.6 | 1.9 | 1.94 | < 5 | 1.92 | 2 | 1.6686 | 1.98 | 1.8543 | 2 | 2.01 | 1.99 | 1.86 | 2.26
| Mean |  | 1.9817 | 2.0878 | 1.4831 | 1.6667 | 2.0167 | 1.9467 |  | 1.925 | 2 | 1.6612 | 1.9717 | 2.0172 | 2.0167 | 1.9517 | 1.935 | 1.855 | 2.3233
| Median |  | 2.005 | 2.0996 | 1.4941 | 1.65 | 2 | 1.94 |  | 1.92 | 2 | 1.6651 | 1.985 | 2.0502 | 2 | 1.95 | 1.935 | 1.855 | 2.31
| Std Dev. |  | 0.093256 | 0.030672 | 0.042088 | 0.08165 | 0.1169 | 0.022509 |  | 0.034496 | 0 | 0.023178 | 0.030605 | 0.13357 | 0.075277 | 0.040208 | 0.034496 | 0.037283 | 0.064704
| Rel.Std.Dev. |  | 4.706% | 1.469% | 2.838% | 4.899% | 5.797% | 1.156% |  | 1.792% | 0% | 1.395% | 1.552% | 6.622% | 3.733% | 2.06% | 1.783% | 2.01% | 2.785%
| PDM3 |  | 0.4299% | 5.811% | -24.84% | -15.53% | 2.204% | -1.344% |  | -2.442% | 1.359% | -15.81% | -0.07689% | 2.23% | 2.204% | -1.09% | -1.935% | -5.99% | 17.75%
| Z-Score (Absolute) |  | 0.24 | 1.54 | 5.84 | 3.6 | 0.67 | 0.18 | 6.57 | 0.45 | NA | 3.67 | 0.12 | 0.68 | 0.67 | 0.12 | 0.33 | 1.3 | 4.41
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 250c (Certified Value 0.29 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
2
| 1 | 1 | 0.29 |  | 0.2725 | < 0.1 | 0.3 | 0.32 | 0.24 | < 5 | 0.3 | 0.3 | 0.3001 | 0.26 | 0.29263 | 0.32 | 0.26 | 0.27 | 0.29 | 0.33 | 1
| 1 | 2 | 0.31 |  | 0.26677 | < 0.1 | 0.3 | 0.34 | 0.24 | < 5 | 0.31 | 0.3 | 0.3121 | 0.26 | 0.25918 | 0.29 | 0.27 | 0.27 | 0.29 | 0.36 | 24
| 1 | 3 | 0.3 |  | 0.28032 | 0.9135 | 0.3 | 0.35 | 0.23 | < 5 | 0.31 | 0.3 | 0.3077 | 0.27 | 0.28067 | 0.3 | 0.28 | 0.28 | 0.29 | 0.34 | 16
| 1 | 4 | 0.3 |  | 0.28267 | 0.1589 | 0.3 | 0.34 | 0.23 | < 5 | 0.3 | 0.3 | 0.3084 | 0.26 | 0.25439 | 0.31 | 0.27 | 0.26 | 0.28 | 0.37 | 0.29075
| 1 | 5 | 0.32 |  | 0.27146 | < 0.1 | 0.3 | 0.32 | 0.23 | < 5 | 0.3 | 0.3 | 0.3011 | 0.27 | 0.29259 | 0.29 | 0.27 | 0.27 | 0.29 | 0.35 | 20
| 1 | 6 | 0.28 |  | 0.27604 | < 0.1 | 0.3 | 0.3 | 0.24 | < 5 | 0.34 | 0.3 | 0.3086 | 0.29 | 0.25137 | 0.3 | 0.28 | 0.26 | 0.31 | 0.34
| Mean |  | 0.3 |  | 0.27496 | 0.5362 | 0.3 | 0.32833 | 0.235 |  | 0.31 | 0.3 | 0.30633 | 0.26833 | 0.27181 | 0.30167 | 0.27167 | 0.26833 | 0.29167 | 0.34833
| Median |  | 0.3 |  | 0.27427 | 0.5362 | 0.3 | 0.33 | 0.235 |  | 0.305 | 0.3 | 0.30805 | 0.265 | 0.26993 | 0.3 | 0.27 | 0.27 | 0.29 | 0.345
| Std Dev. |  | 0.014142 |  | 0.0059086 | 0.53358 | 0 | 0.018348 | 0.0054772 |  | 0.015492 | 0 | 0.0047069 | 0.01169 | 0.019104 | 0.01169 | 0.0075277 | 0.0075277 | 0.0098319 | 0.01472
| Rel.Std.Dev. |  | 4.714% |  | 2.149% | 99.51% | 0% | 5.588% | 2.331% |  | 4.997% | 0% | 1.537% | 4.357% | 7.028% | 3.875% | 2.771% | 2.805% | 3.371% | 4.226%
| PDM3 |  | 3.183% |  | -5.43% | 84.42% | 3.183% | 12.93% | -19.17% |  | 6.623% | 3.183% | 5.362% | -7.708% | -6.514% | 3.756% | -6.562% | -7.708% | 0.317% | 19.81%
| Z-Score (Absolute) |  | 0.12 | NA | 0.58 | 2.34 | 0.12 | 0.91 | 1.7 | 61.68 | 0.4 | NA | 0.29 | 0.77 | 0.67 | 0.16 | 0.68 | 0.77 | 0.12 | 1.47
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 250c (Certified Value 1.72 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.7406 | 1.7354 | 1.7133 | 2.4179 | 1.68 | 1.61 | 1.67 | 1.72 | 1.65 | 1.7 | 1.8134 | 1.82 | 1.7144 | 1.73 | 1.7 | 1.74 | 1.69 | 1.576 | 1
| 1 | 2 | 1.6966 | 1.7296 | 1.7223 | 2.4218 | 1.68 | 1.77 | 1.69 | 1.66 | 1.66 | 1.76 | 1.7362 | 1.76 | 1.745 | 1.71 | 1.7 | 1.77 | 1.69 | 1.614 | #N/A
| 1 | 3 | 1.6982 | 1.7052 | 1.7259 | 2.3483 | 1.67 | 1.8 | 1.65 | 1.69 | 1.74 | 1.75 | 1.8302 | 1.77 | 1.6795 | 1.71 | 1.71 | 1.71 | 1.67 | 1.608 | 16
| 1 | 4 | 1.7398 | 1.7316 | 1.7146 | 2.3888 | 1.66 | 1.65 | 1.66 | 1.74 | 1.72 | 1.79 | 1.8181 | 1.78 | 1.7151 | 1.68 | 1.76 | 1.76 | 1.7 | 1.629 | 1.7157
| 1 | 5 | 1.7711 | 1.7575 | 1.7329 | 2.4111 | 1.67 | 1.61 | 1.65 | 1.7 | 1.7 | 1.81 | 1.822 | 1.83 | 1.6911 | 1.68 | 1.72 | 1.75 | 1.79 | 1.596 | 21
| 1 | 6 | 1.6465 | 1.7889 | 1.7539 | 2.4158 | 1.67 | 1.54 | 1.69 | 1.74 | 1.7 | 1.8 | 1.8258 | 1.82 | 1.7514 | 1.71 | 1.74 | 1.78 | 1.72 | 1.596
| Mean |  | 1.7155 | 1.7414 | 1.7271 | 2.4006 | 1.6717 | 1.6633 | 1.6683 | 1.7083 | 1.695 | 1.7683 | 1.8076 | 1.7967 | 1.7161 | 1.7033 | 1.7217 | 1.7517 | 1.71 | 1.6032
| Median |  | 1.719 | 1.7335 | 1.7241 | 2.4135 | 1.67 | 1.63 | 1.665 | 1.71 | 1.7 | 1.775 | 1.82 | 1.8 | 1.7147 | 1.71 | 1.715 | 1.755 | 1.695 | 1.602
| Std Dev. |  | 0.044112 | 0.028623 | 0.014991 | 0.028177 | 0.0075277 | 0.10113 | 0.018348 | 0.031252 | 0.034496 | 0.040702 | 0.035469 | 0.030111 | 0.028461 | 0.019664 | 0.024014 | 0.024833 | 0.042426 | 0.018159
| Rel.Std.Dev. |  | 2.571% | 1.644% | 0.868% | 1.174% | 0.4503% | 6.08% | 1.1% | 1.829% | 2.035% | 2.302% | 1.962% | 1.676% | 1.659% | 1.154% | 1.395% | 1.418% | 2.481% | 1.133%
| PDM3 |  | -0.01541% | 1.495% | 0.6644% | 39.92% | -2.568% | -3.054% | -2.763% | -0.4312% | -1.208% | 3.066% | 5.355% | 4.717% | 0.01977% | -0.7226% | 0.3459% | 2.094% | -0.334% | -6.561%
| Z-Score (Absolute) |  | 0.01 | 0.56 | 0.25 | 15.02 | 0.97 | 1.15 | 1.04 | 0.16 | 0.46 | 1.15 | 2.01 | 1.77 | 0.01 | 0.27 | 0.13 | 0.79 | 0.13 | 2.47
| Analytical results for Cd in OREAS 250c (Certified Value 0.075 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.09 | < 0.2 | 0.532 | 0.08 | < 0.3 | 0.2 | < 0.5 | 0.06 | 0.1 | 0.0899 | 0.07 | 0.086222 | 0.32 | 0.05 | 0.07 | 0.11 |  |  | 1
| 1 | 2 | 0.1 | < 0.2 | 0.6606 | 0.1 | < 0.3 | 0.2 | < 0.5 | 0.07 | 0.2 | 0.0875 | 0.05 | 0.079847 | 0.29 | 0.06 | 0.07 | 0.05 |  |  | 25
| 1 | 3 | 0.09 | < 0.2 | 0.5037 | 0.1 | < 0.3 | 0.22 | < 0.5 | 0.07 | 0.2 | 0.0894 | 0.06 | 0.05898 | 0.29 | 0.07 | 0.07 | 0.05 |  |  | 16
| 1 | 4 | 0.09 | < 0.2 | 0.4533 | 0.09 | < 0.3 | 0.2 | < 0.5 | 0.08 | 0.2 | 0.0985 | 0.05 | < 0.05 | 0.3 | 0.06 | 0.06 | 0.07 |  |  | 0.074604
| 1 | 5 | 0.08 | < 0.2 | 0.4546 | 0.06 | < 0.3 | 0.21 | < 0.5 | 0.09 | 0.2 | 0.0956 | 0.04 | 0.084289 | 0.29 | 0.07 | 0.07 | 0.06 |  |  | 22
| 1 | 6 | 0.08 | < 0.2 | 0.6039 | 0.1 | < 0.3 | 0.21 | < 0.5 | 0.03 | 0.1 | 0.0937 | 0.06 | < 0.05 | 0.3 | 0.07 | 0.08 | 0.08
| Mean |  | 0.088333 |  | 0.53468 | 0.088333 |  | 0.20667 |  | 0.066667 | 0.16667 | 0.092433 | 0.055 | 0.077335 | 0.29833 | 0.063333 | 0.07 | 0.07
| Median |  | 0.09 |  | 0.51785 | 0.095 |  | 0.205 |  | 0.07 | 0.2 | 0.0918 | 0.055 | 0.082068 | 0.295 | 0.065 | 0.07 | 0.065
| Std Dev. |  | 0.0075277 |  | 0.083246 | 0.016021 |  | 0.008165 |  | 0.020656 | 0.05164 | 0.0042046 | 0.010488 | 0.012524 | 0.01169 | 0.008165 | 0.0063246 | 0.022804
| Rel.Std.Dev. |  | 8.522% |  | 15.57% | 18.14% |  | 3.951% |  | 30.98% | 30.98% | 4.549% | 19.07% | 16.19% | 3.919% | 12.89% | 9.035% | 32.58%
| PDM3 |  | 18.4% |  | 616.7% | 18.4% |  | 177% |  | -10.64% | 123.4% | 23.9% | -26.28% | 3.66% | 299.9% | -15.11% | -6.171% | -6.171%
| Z-Score (Absolute) |  | 0 | 0.31 | 12.04 | 0 | 1.66 | 3.19 | 4.36 | 0.58 | NA | 0.11 | 0.9 | 0.77 | 5.66 | 0.67 | 0.49 | 0.49
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 250c (Certified Value 70 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
1
| 1 | 1 | 72.38 | 72.118 | 67.7889 | 67.8 | 60 | 69 | 64 | 58.4957 | 74.1 | 74 | 72.3 | 71.8 | 70.2 | 70.13 |  |  |  |  | 1
| 1 | 2 | 74.07 | 73.8003 | 66.9582 | 73.2 | 57 | 68 | 65 | 56.982 | 69.2 | 73.2 | 72.1 | 73.5 | 69.4 | 68.14 |  |  |  |  | 26
| 1 | 3 | 73.18 | 72.2079 | 67.4111 | 73.5 | 59 | 71.9 | 66 | 59.1238 | 70.7 | 72.2 | 71.7 | 72 | 69.5 | 71.11 |  |  |  |  | 16
| 1 | 4 | 74.05 | 75.3909 | 68.055 | 68.3 | 56 | 69.7 | 65 | 59.4795 | 66.2 | 74.2 | 74.1 | 71.6 | 69.6 | 70.3 |  |  |  |  | 70.4375
| 1 | 5 | 72.17 | 75.0649 | 63.9837 | 67.2 | 61 | 71 | 67 | 55.8135 | 67.9 | 70.8 | 73.4 | 69.8 | 72.3 | 67.3 |  |  |  |  | 23
| 1 | 6 | 69.54 | 74.5761 | 65.6162 | 67.2 | 63 | 69.6 | 64 | 57.829 | 72.1 | 73.3 | 76.4 | 72.5 | 71.4 | 67.26
| Mean |  | 72.565 | 73.8597 | 66.6355 | 69.5333 | 59.3333 | 69.8667 | 65.1667 | 57.9539 | 70.0333 | 72.95 | 73.3333 | 71.8667 | 70.4 | 69.04
| Median |  | 72.78 | 74.1882 | 67.1846 | 68.05 | 59.5 | 69.65 | 65 | 58.1624 | 69.95 | 73.25 | 72.85 | 71.9 | 69.9 | 69.135
| Std Dev. |  | 1.6846 | 1.4196 | 1.5571 | 2.9864 | 2.582 | 1.3967 | 1.169 | 1.3807 | 2.8703 | 1.2677 | 1.7466 | 1.2193 | 1.1916 | 1.6773
| Rel.Std.Dev. |  | 2.322% | 1.922% | 2.337% | 4.295% | 4.352% | 1.999% | 1.794% | 2.382% | 4.098% | 1.738% | 2.382% | 1.697% | 1.693% | 2.429%
| PDM3 |  | 3.02% | 4.858% | -5.398% | -1.284% | -15.76% | -0.8104% | -7.483% | -17.72% | -0.5738% | 3.567% | 4.111% | 2.029% | -0.05326% | -1.984%
| Z-Score (Absolute) |  | 0.63 | 0.94 | 0.8 | 0.1 | 2.55 | 0.02 | 1.15 | 2.88 | 0.02 | 0.72 | 0.81 | 0.46 | 0.11 | 0.22
| Analytical results for Co in OREAS 250c (Certified Value 20.6 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 20.2 | 20.43 | 21.2858 | 21.6533 | 20 | 20 | 20.1 | 23.5615 | 19.6 | 19.7 | 18.5014 | 19.9 | 19.8288 | 20.84 | 20.6 | 21.7 | 19.4 | 21.4 | 1
| 1 | 2 | 20.3 | 21.402 | 20.9166 | 21.8105 | 20 | 20 | 20.6 | 22.6518 | 19 | 19.9 | 17.9132 | 19.3 | 20.0426 | 20.56 | 21.6 | 21.8 | 19.2 | 20.9 | 27
| 1 | 3 | 20.6 | 21.609 | 20.7864 | 21.6408 | 20 | 20 | 20.6 | 23.3845 | 20.2 | 19.2 | 18.0393 | 19.6 | 19.4435 | 20.6 | 21 | 21.9 | 19.4 | 21.4 | 16
| 1 | 4 | 20.6 | 21.276 | 21.145 | 23.2004 | 19 | 21 | 19.8 | 24.0084 | 19.8 | 19.5 | 18.4842 | 17.9 | 19.7809 | 20.59 | 21.7 | 21.8 | 19.6 | 21.2 | 20.5756
| 1 | 5 | 20.1 | 23.454 | 21.0347 | 22.3453 | 20 | 20 | 20.7 | 23.0113 | 19.6 | 19.7 | 17.5513 | 18.7 | 20.4717 | 20.04 | 21.9 | 20.8 | 20.1 | 20.4 | 24
| 1 | 6 | 20.1 | 22.536 | 20.9493 | 23.785 | 20 | 21 | 20.6 | 23.8524 | 19.4 | 19.4 | 18.5548 | 19.5 | 20.4736 | 20.6 | 21.6 | 21.8 | 19.9 | 20.1
| Mean |  | 20.3167 | 21.7845 | 21.0196 | 22.4059 | 19.8333 | 20.3333 | 20.4 | 23.4117 | 19.6 | 19.5667 | 18.174 | 19.15 | 20.0069 | 20.5383 | 21.4 | 21.6333 | 19.6 | 20.9
| Median |  | 20.25 | 21.5055 | 20.992 | 22.0779 | 20 | 20 | 20.6 | 23.473 | 19.6 | 19.6 | 18.2617 | 19.4 | 19.9357 | 20.595 | 21.6 | 21.8 | 19.5 | 21.05
| Std Dev. |  | 0.23166 | 1.0604 | 0.17707 | 0.89923 | 0.40825 | 0.5164 | 0.36332 | 0.51211 | 0.4 | 0.25033 | 0.40553 | 0.73144 | 0.40875 | 0.26461 | 0.49396 | 0.41312 | 0.34059 | 0.54406
| Rel.Std.Dev. |  | 1.14% | 4.868% | 0.8424% | 4.013% | 2.058% | 2.54% | 1.781% | 2.187% | 2.041% | 1.279% | 2.231% | 3.82% | 2.043% | 1.288% | 2.308% | 1.91% | 1.738% | 2.603%
| PDM3 |  | -1.258% | 5.875% | 2.158% | 8.895% | -3.608% | -1.177% | -0.8535% | 13.78% | -4.742% | -4.904% | -11.67% | -6.929% | -2.764% | -0.1812% | 4.007% | 5.141% | -4.742% | 1.577%
| Z-Score (Absolute) |  | 0.04 | 1.25 | 0.57 | 1.79 | 0.47 | 0.03 | 0.03 | 2.68 | 0.67 | 0.7 | 1.93 | 1.07 | 0.32 | 0.15 | 0.91 | 1.11 | 0.67 | 0.47
| Analytical results for Cr in OREAS 250c (Certified Value 100 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 107 | 101.6375 | 113.6058 | 104 | 96 | 96 | 94.7578 | 97 | 85 | 83.9163 | 109 | 97.7566 | 104 | 103 | 100 | 97 | 95 |  | 1
| 1 | 2 | 111 | 98.1263 | 107.6476 | 109 | 83 | 98 | 91.7965 | 98 | 87 | 78.8248 | 105 | 103.4148 | 101 | 104 | 102 | 96 | 95 |  | 28
| 1 | 3 | 115 | 100.6518 | 119.9893 | 107 | 84 | 96 | 87.5393 | 102 | 89 | 81.1228 | 109 | 98.1177 | 102 | 105 | 103 | 98 | 96 |  | 16
| 1 | 4 | 112 | 100.961 | 125.1497 | 103 | 81 | 97 | 95.4295 | 101 | 91 | 84.116 | 105 | 97.363 | 103 | 104 | 104 | 94 | 98 |  | 100.2348
| 1 | 5 | 114 | 102.3369 | 127.9901 | 104 | 82 | 94 | 92.6864 | 100 | 88 | 78.0284 | 108 | 103.2774 | 100 | 102 | 99 | 101 | 96 |  | 25
| 1 | 6 | 113 | 102.388 | 133.0033 | 101 | 75 | 96 | 93.6502 | 100 | 113 | 82.5179 | 110 | 101.8284 | 102 | 104 | 105 | 98 | 96
| Mean |  | 112 | 101.0169 | 121.231 | 104.6667 | 83.5 | 96.1667 | 92.6433 | 99.6667 | 92.1667 | 81.421 | 107.6667 | 100.293 | 102 | 103.6667 | 102.1667 | 97.3333 | 96
| Median |  | 112.5 | 101.2992 | 122.5695 | 104 | 82.5 | 96 | 93.1683 | 100 | 88.5 | 81.8204 | 108.5 | 99.9731 | 102 | 104 | 102.5 | 97.5 | 96
| Std Dev. |  | 2.8284 | 1.5811 | 9.4215 | 2.8752 | 6.892 | 1.3292 | 2.8291 | 1.8619 | 10.4003 | 2.5703 | 2.1602 | 2.8552 | 1.4142 | 1.0328 | 2.3166 | 2.3381 | 1.0954
| Rel.Std.Dev. |  | 2.525% | 1.565% | 7.772% | 2.747% | 8.254% | 1.382% | 3.054% | 1.868% | 11.28% | 3.157% | 2.006% | 2.847% | 1.386% | 0.9963% | 2.267% | 2.402% | 1.141%
| PDM3 |  | 11.74% | 0.7803% | 20.95% | 4.422% | -16.7% | -4.059% | -7.574% | -0.5668% | -8.049% | -18.77% | 7.415% | 0.05812% | 1.761% | 3.424% | 1.927% | -2.895% | -4.225%
| Z-Score (Absolute) |  | 1.84 | 0.11 | 3.29 | 0.69 | 2.64 | 0.65 | 1.2 | 0.1 | 1.28 | 2.96 | 1.16 | 0 | 0.27 | 0.53 | 0.29 | 0.46 | 0.67
| Analytical results for Cs in OREAS 250c (Certified Value 2.62 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.66 | 2.6267 | 2.6 | 2.43 | 2.6 | 2.58 | 2.575 | 2.65 | 2.6411 | 2.57 | 2.51 | 2.75 | 2.69 | 2.79 |  |  |  |  | 1
| 1 | 2 | 2.7 | 2.7346 | 2.5 | 2.61 | 2.71 | 2.58 | 2.5726 | 2.57 | 2.541 | 2.54 | 2.54 | 2.75 | 2.61 | 2.78 |  |  |  |  | 29
| 1 | 3 | 2.71 | 2.8227 | 2.6 | 2.65 | 2.63 | 2.69 | 2.6267 | 2.64 | 2.562 | 2.52 | 2.53 | 2.74 | 2.63 | 2.82 |  |  |  |  | 16
| 1 | 4 | 2.75 | 2.6516 | 2.5 | 2.4 | 2.57 | 2.65 | 2.6836 | 2.45 | 2.3357 | 2.56 | 2.63 | 2.75 | 2.59 | 2.88 |  |  |  |  | 2.6165
| 1 | 5 | 2.62 | 2.7301 | 2.5 | 2.39 | 2.64 | 2.61 | 2.5447 | 2.5 | 2.7075 | 2.48 | 2.6 | 2.92 | 2.74 | 2.71 |  |  |  |  | 26
| 1 | 6 | 2.49 | 2.6509 | 2.5 | 2.39 | 2.73 | 2.61 | 2.6332 | 2.64 | 2.3149 | 2.54 | 2.64 | 2.77 | 2.69 | 2.68
| Mean |  | 2.655 | 2.7028 | 2.5333 | 2.4783 | 2.6467 | 2.62 | 2.606 | 2.575 | 2.517 | 2.535 | 2.575 | 2.78 | 2.6583 | 2.7767
| Median |  | 2.68 | 2.6909 | 2.5 | 2.415 | 2.635 | 2.61 | 2.6009 | 2.605 | 2.5515 | 2.54 | 2.57 | 2.75 | 2.66 | 2.785
| Std Dev. |  | 0.09225 | 0.073804 | 0.05164 | 0.11907 | 0.062183 | 0.042895 | 0.051038 | 0.084083 | 0.15999 | 0.032094 | 0.055408 | 0.069282 | 0.057417 | 0.072847
| Rel.Std.Dev. |  | 3.475% | 2.731% | 2.038% | 4.804% | 2.349% | 1.637% | 1.959% | 3.265% | 6.356% | 1.266% | 2.152% | 2.492% | 2.16% | 2.624%
| PDM3 |  | 1.471% | 3.297% | -3.179% | -5.281% | 1.153% | 0.1334% | -0.4029% | -1.586% | -3.801% | -3.115% | -1.586% | 6.248% | 1.598% | 6.121%
| Z-Score (Absolute) |  | 0.46 | 0.98 | 0.87 | 1.47 | 0.37 | 0.08 | 0.08 | 0.42 | 1.05 | 0.85 | 0.42 | 1.83 | 0.5 | 1.79
| Analytical results for Cu in OREAS 250c (Certified Value 24.1 ppm) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 26 | 24.4677 | 22.889 | 22 | 24 | 28 | 21.9 | 23.7 | 25.6 | 16.0321 | 32.6 | 23.4771 | 24.5 | 23.8 | 26.4 | 24.3 | 24 |  | 1
| 1 | 2 | 25.8 | 23.992 | 22.5472 | 22 | 25 | 29 | 22.3067 | 22.4 | 25.8 | 13.7828 | 26.4 | 22.9819 | 24.4 | 24.2 | 26.9 | 24.3 | 24 |  | 30
| 1 | 3 | 25.4 | 24.3378 | 22.912 | 22 | 24 | 29 | 20.8133 | 24.2 | 25.9 | 15.0682 | 23.9 | 18.3365 | 24.5 | 23.7 | 26.9 | 24.2 | 24.2 |  | 16
| 1 | 4 | 26.4 | 23.8973 | 23.2451 | 22 | 25 | 28 | 20.6833 | 23.5 | 25.9 | 16.6097 | 21.5 | 18.6438 | 25.1 | 24.6 | 26.9 | 24 | 24.7 |  | 24.1278
| 1 | 5 | 25.2 | 24.0855 | 22.796 | 22 | 24 | 29 | 20.1633 | 23.2 | 26.2 | 13.6287 | 22.8 | 17.6485 | 24.6 | 24 | 26 | 25.3 | 24.5 |  | 27
| 1 | 6 | 24.2 | 24.1289 | 23.211 | 22 | 24 | 29 | 21.2233 | 22.7 | 26.2 | 15.6036 | 25.3 | 19.0991 | 24.3 | 24.7 | 26.1 | 25.2 | 24
| Mean |  | 25.5 | 24.1515 | 22.9334 | 22 | 24.3333 | 28.6667 | 21.1817 | 23.2833 | 25.9333 | 15.1209 | 25.4167 | 20.0312 | 24.5667 | 24.1667 | 26.5333 | 24.55 | 24.2333
| Median |  | 25.6 | 24.1072 | 22.9005 | 22 | 24 | 29 | 21.0183 | 23.35 | 25.9 | 15.3359 | 24.6 | 18.8714 | 24.5 | 24.1 | 26.65 | 24.3 | 24.1
| Std Dev. |  | 0.76681 | 0.21432 | 0.26256 | 0 | 0.5164 | 0.5164 | 0.80055 | 0.66156 | 0.23381 | 1.2083 | 3.926 | 2.5268 | 0.28048 | 0.41312 | 0.42269 | 0.55408 | 0.30111
| Rel.Std.Dev. |  | 3.007% | 0.8874% | 1.145% | 0% | 2.122% | 1.801% | 3.779% | 2.841% | 0.9016% | 7.991% | 15.45% | 12.61% | 1.142% | 1.709% | 1.593% | 2.257% | 1.243%
| PDM3 |  | 5.687% | 0.09846% | -4.95% | -8.819% | 0.852% | 18.81% | -12.21% | -3.5% | 7.483% | -37.33% | 5.342% | -16.98% | 1.819% | 0.1612% | 9.97% | 1.75% | 0.4375%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 0.69 | 1.19 | 0.05 | 2.36 | 1.62 | 0.51 | 0.9 | 4.85 | 0.63 | 2.24 | 0.18 | 0.04 | 1.22 | 0.17 | 0
| Analytical results for Dy in OREAS 250c (Certified Value 3.67 ppm) | [19 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 3.58 | 3.7373 | 3.6835 | 3.6 | 3.43 | 2.8185 | 3.51 | [11 empty] | 1
| 1 | 2 | 3.81 | 3.7469 | 4.0352 | 3.9 | 3.53 | 2.7014 | 3.38 | [11 empty] | 4
| 1 | 3 | 3.74 | 3.6457 | 4.079 | 3.8 | 3.44 | 2.7112 | 3.35 | [11 empty] | 16
| 1 | 4 | 3.83 | 3.8694 | 3.897 | 3.7 | 3.44 | 2.8465 | 3.47 | [11 empty] | 3.6654
| 1 | 5 | 3.76 | 3.7359 | 3.9429 | 3.5 | 3.51 | 2.6362 | 3.31 | [11 empty] | 28
| 1 | 6 | 3.57 | 3.8781 | 3.9153 | 3.7 | 3.56 | 2.7778 | 3.37
| Mean |  | 3.715 | 3.7689 | 3.9255 | 3.7 | 3.485 | 2.7486 | 3.3983
| Median |  | 3.75 | 3.7421 | 3.9291 | 3.7 | 3.475 | 2.7445 | 3.375
| Std Dev. |  | 0.11327 | 0.089179 | 0.1381 | 0.14142 | 0.055408 | 0.079479 | 0.076004
| Rel.Std.Dev. |  | 3.049% | 2.366% | 3.518% | 3.822% | 1.59% | 2.892% | 2.237%
| PDM3 |  | 1.352% | 2.822% | 7.094% | 0.9426% | -4.923% | -25.01% | -7.287%
| Z-Score (Absolute) |  | 0.05 | 0.22 | 0.71 | 0 | 0.67 | 2.98 | 0.95
| Analytical results for Er in OREAS 250c (Certified Value 1.59 ppm) | [19 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.6 | 1.6667 | 1.4978 | 1.7 | 1.54 | 3.3635 | 1.47 | [11 empty] | 1
| 1 | 2 | 1.64 | 1.7336 | 1.5909 | 1.7 | 1.53 | 3.2389 | 1.46 | [11 empty] | 5
| 1 | 3 | 1.68 | 1.6801 | 1.5638 | 1.7 | 1.54 | 3.2125 | 1.46 | [11 empty] | 16
| 1 | 4 | 1.68 | 1.7701 | 1.5394 | 1.7 | 1.5 | 3.3911 | 1.48 | [11 empty] | 1.5937
| 1 | 5 | 1.63 | 1.7621 | 1.5702 | 1.6 | 1.56 | 3.1633 | 1.37 | [11 empty] | 29
| 1 | 6 | 1.52 | 1.7615 | 1.5381 | 1.7 | 1.57 | 3.3621 | 1.37
| Mean |  | 1.625 | 1.729 | 1.55 | 1.6833 | 1.54 | 3.2886 | 1.435
| Median |  | 1.635 | 1.7476 | 1.5516 | 1.7 | 1.54 | 3.3005 | 1.46
| Std Dev. |  | 0.059917 | 0.045042 | 0.032418 | 0.040825 | 0.024495 | 0.095369 | 0.050892
| Rel.Std.Dev. |  | 3.687% | 2.605% | 2.091% | 2.425% | 1.591% | 2.9% | 3.546%
| PDM3 |  | 1.962% | 8.489% | -2.742% | 5.622% | -3.371% | 106.3% | -9.96%
| Z-Score (Absolute) |  | 0 | 0.83 | 0.59 | 0.46 | 0.67 | 13.2 | 1.51
| Analytical results for Eu in OREAS 250c (Certified Value 1.6 ppm) | [19 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.6 | 0.59135 | 1.5514 | 1.5 | 1.59 | 1.4223 | 1.67 | [11 empty] | 1
| 1 | 2 | 1.68 | 0.61607 | 1.5564 | 1.63 | 1.63 | 1.3871 | 1.63 | [11 empty] | 6
| 1 | 3 | 1.69 | 0.5578 | 1.6352 | 1.66 | 1.59 | 1.3859 | 1.64 | [11 empty] | 16
| 1 | 4 | 1.75 | 0.61535 | 1.6142 | 1.56 | 1.57 | 1.4315 | 1.56 | [11 empty] | 1.6023
| 1 | 5 | 1.64 | 0.58987 | 1.6622 | 1.49 | 1.61 | 1.3512 | 1.61 | [11 empty] | 30
| 1 | 6 | 1.55 | 0.57441 | 1.5196 | 1.52 | 1.6 | 1.4216 | 1.56
| Mean |  | 1.6517 | 0.59081 | 1.5898 | 1.56 | 1.5983 | 1.3999 | 1.6117
| Median |  | 1.66 | 0.59061 | 1.5853 | 1.54 | 1.595 | 1.4044 | 1.62
| Std Dev. |  | 0.070828 | 0.022806 | 0.055522 | 0.070711 | 0.020412 | 0.030674 | 0.04446
| Rel.Std.Dev. |  | 4.288% | 3.86% | 3.492% | 4.533% | 1.277% | 2.191% | 2.759%
| PDM3 |  | 3.081% | -63.13% | -0.7778% | -2.64% | -0.2476% | -12.63% | 0.5845%
| Z-Score (Absolute) |  | 1.4 | 22.58 | 0 | 0.67 | 0.19 | 4.29 | 0.49
| Analytical results for Fe in OREAS 250c (Certified Value 4.04 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.13 | 4.2462 | 4.1999 | 3.2688 | 4.13 | 3.64 | 3.95 | 4 | 3.98 | 3.77 | 4.0957 | 4.17 | 4.1122 | 4.09 | 3.97 | 3.99 | 4 | 3.82 | 1
| 1 | 2 | 4.02 | 4.3044 | 4.1986 | 3.2867 | 4.2 | 4.02 | 3.99 | 3.89 | 3.94 | 3.88 | 4.0163 | 4.05 | 4.164 | 4.05 | 4 | 4.12 | 4.01 | 3.81 | #N/A
| 1 | 3 | 4.02 | 4.2917 | 4.2137 | 3.229 | 4.11 | 4.11 | 3.93 | 3.95 | 4.17 | 3.83 | 4.0607 | 4.12 | 4.0636 | 4.07 | 3.99 | 3.93 | 3.93 | 3.77 | 16
| 1 | 4 | 4.11 | 4.2952 | 4.1532 | 3.2503 | 4.15 | 3.76 | 3.93 | 3.99 | 4.09 | 3.87 | 4.1411 | 4.06 | 4.1732 | 4.05 | 4.11 | 4.06 | 3.92 | 3.9 | 4.0366
| 1 | 5 | 4.17 | 4.2953 | 4.1699 | 3.3109 | 4.06 | 3.65 | 4 | 3.97 | 4.02 | 3.91 | 3.9978 | 4.16 | 4.0528 | 4.02 | 4.09 | 4.05 | 4.15 | 3.8 | 31
| 1 | 6 | 3.88 | 4.2108 | 4.1989 | 3.3116 | 4.13 | 3.53 | 3.99 | 4.09 | 4.06 | 3.9 | 4.1416 | 4.19 | 4.1372 | 4.08 | 4.13 | 4.1 | 4.04 | 3.77
| Mean |  | 4.055 | 4.2739 | 4.1891 | 3.2762 | 4.13 | 3.785 | 3.965 | 3.9817 | 4.0433 | 3.86 | 4.0756 | 4.125 | 4.1172 | 4.06 | 4.0483 | 4.0417 | 4.0083 | 3.8117
| Median |  | 4.065 | 4.2934 | 4.1988 | 3.2777 | 4.13 | 3.705 | 3.97 | 3.98 | 4.04 | 3.875 | 4.0782 | 4.14 | 4.1247 | 4.06 | 4.045 | 4.055 | 4.005 | 3.805
| Std Dev. |  | 0.10483 | 0.037159 | 0.022649 | 0.033204 | 0.046043 | 0.23054 | 0.032094 | 0.065853 | 0.082138 | 0.052154 | 0.061375 | 0.058907 | 0.050516 | 0.025298 | 0.069402 | 0.070828 | 0.083766 | 0.047924
| Rel.Std.Dev. |  | 2.585% | 0.8694% | 0.5407% | 1.013% | 1.115% | 6.091% | 0.8094% | 1.654% | 2.031% | 1.351% | 1.506% | 1.428% | 1.227% | 0.6231% | 1.714% | 1.752% | 2.09% | 1.257%
| PDM3 |  | 0.4565% | 5.881% | 3.777% | -18.84% | 2.315% | -6.232% | -1.773% | -1.36% | 0.1675% | -4.374% | 0.9658% | 2.191% | 1.997% | 0.5804% | 0.2914% | 0.1262% | -0.6996% | -5.572%
| Z-Score (Absolute) |  | 0.08 | 2.04 | 1.28 | 6.9 | 0.75 | 2.34 | 0.72 | 0.57 | 0.02 | 1.67 | 0.27 | 0.71 | 0.64 | 0.13 | 0.02 | 0.04 | 0.34 | 2.1
| Analytical results for Ga in OREAS 250c (Certified Value 15.3 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 15.75 | 13.56 | 15.326 | 16.9 | 8.9 | 16.57 | 16.9459 | 15.15 | 12.8284 | 14.95 | 15.7176 | 16.17 | 14.35 | 14.95 | 15.1 | 16.73 |  |  | 1
| 1 | 2 | 16.27 | 14.18 | 15.6293 | 16.8 | 10.2 | 16.68 | 16.3906 | 14.85 | 12.6577 | 14.5 | 14.7223 | 16.06 | 14.65 | 15.3 | 14.9 | 16.46 |  |  | 34
| 1 | 3 | 16.1 | 14.25 | 15.5914 | 16.2 | 10.8 | 16.62 | 15.2611 | 15.7 | 12.5931 | 14.95 | 14.9755 | 16.07 | 14.4 | 15.35 | 14.8 | 16.26 |  |  | 16
| 1 | 4 | 16.11 | 14.04 | 15.0284 | 16.5 | 9.9 | 16.34 | 17.1736 | 15.25 | 12.9794 | 13.55 | 13.7832 | 16.17 | 15.05 | 15.7 | 14.5 | 16.94 |  |  | 15.3212
| 1 | 5 | 15.44 | 15.42 | 15.3969 | 16.4 | 9.9 | 16.67 | 16.6688 | 15.15 | 12.4485 | 14.3 | 15.5275 | 15.58 | 14.75 | 14.95 | 15.75 | 15.97 |  |  | 32
| 1 | 6 | 14.94 | 14.52 | 15.3318 | 16.4 | 8.8 | 16.56 | 17.4301 | 15.15 | 13.0814 | 14.65 | 13.5151 | 16.04 | 15.1 | 15.4 | 15.1 | 16.05
| Mean |  | 15.7683 | 14.3283 | 15.384 | 16.5333 | 9.75 | 16.5733 | 16.645 | 15.2083 | 12.7648 | 14.4833 | 14.7069 | 16.015 | 14.7167 | 15.275 | 15.025 | 16.4017
| Median |  | 15.925 | 14.215 | 15.3643 | 16.45 | 9.9 | 16.595 | 16.8073 | 15.15 | 12.7431 | 14.575 | 14.8489 | 16.065 | 14.7 | 15.325 | 15 | 16.36
| Std Dev. |  | 0.50452 | 0.6213 | 0.21721 | 0.26583 | 0.77136 | 0.12453 | 0.7703 | 0.27644 | 0.24138 | 0.52313 | 0.89898 | 0.22043 | 0.3157 | 0.28766 | 0.41923 | 0.38238
| Rel.Std.Dev. |  | 3.2% | 4.336% | 1.412% | 1.608% | 7.911% | 0.7514% | 4.628% | 1.818% | 1.891% | 3.612% | 6.113% | 1.376% | 2.145% | 1.883% | 2.79% | 2.331%
| PDM3 |  | 2.919% | -6.48% | 0.4099% | 7.912% | -36.36% | 8.173% | 8.641% | -0.7365% | -16.69% | -5.469% | -4.01% | 4.529% | -3.946% | -0.3014% | -1.933% | 7.052%
| Z-Score (Absolute) |  | 0.46 | 0.8 | 0.13 | 1.14 | 4.84 | 1.17 | 1.24 | 0.03 | 2.18 | 0.67 | 0.47 | 0.68 | 0.46 | 0.03 | 0.19 | 1.02
| Analytical results for Gd in OREAS 250c (Certified Value 5.2 ppm) | [19 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 5.12 | 6.0959 | 5.1204 | 5.2 | 5.05 | 3.0699 | 5.25 | [11 empty] | 1
| 1 | 2 | 5.06 | 6.1232 | 5.5961 | 5.6 | 5.1 | 2.93 | 5.09 | [11 empty] | 7
| 1 | 3 | 4.91 | 5.9986 | 5.6524 | 5.7 | 5.06 | 2.845 | 5.18 | [11 empty] | 16
| 1 | 4 | 5.01 | 6.3408 | 5.2505 | 5.4 | 4.92 | 2.992 | 5.29 | [11 empty] | 5.2047
| 1 | 5 | 4.79 | 6.3386 | 5.6373 | 5 | 5.1 | 2.7791 | 5.27 | [11 empty] | 33
| 1 | 6 | 4.64 | 6.3255 | 5.4295 | 5.3 | 5.11 | 2.9859 | 5.14
| Mean |  | 4.9217 | 6.2038 | 5.4477 | 5.3667 | 5.0567 | 2.9337 | 5.2033
| Median |  | 4.96 | 6.2244 | 5.5128 | 5.35 | 5.08 | 2.958 | 5.215
| Std Dev. |  | 0.1806 | 0.14965 | 0.22192 | 0.2582 | 0.071181 | 0.10627 | 0.079415
| Rel.Std.Dev. |  | 3.67% | 2.412% | 4.074% | 4.811% | 1.408% | 3.623% | 1.526%
| PDM3 |  | -5.438% | 19.2% | 4.67% | 3.112% | -2.844% | -43.63% | -0.02585%
| Z-Score (Absolute) |  | 0.78 | 2.76 | 0.67 | 0.45 | 0.4 | 6.26 | 0
| Analytical results for Ge in OREAS 250c (Indicative Value 0.14 ppm) | [19 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.1 | < 0.2 | 1.4 | < 0.1 | 0.08 | 0.21 | < 0.05 | 0.0332 | 0.12 | 0.17134 | 0.12 | 0.22 | 0.08 | [5 empty] | 1
| 1 | 2 | 1.2 | < 0.2 | 1.5 | < 0.1 | 0.09 | 0.21 | 0.17 | 0.1689 | 0.1 | 0.17751 | 0.15 | 0.24 | 0.11 | [5 empty] | 1
| 1 | 3 | 1.3 | < 0.2 | 1.4 | < 0.1 | 0.08 | 0.21 | < 0.05 | 0.1057 | 0.11 | 0.15833 | 0.13 | 0.23 | 0.12 | [5 empty] | 16
| 1 | 4 | 1.2 | < 0.2 | 1.6 | < 0.1 | 0.08 | 0.25 | 0.11 | 0.0444 | 0.1 | 0.15875 | 0.14 | 0.23 | 0.09 | [5 empty] | 0.14121
| 1 | 5 | 1.2 | < 0.2 | 1.4 | < 0.1 | 0.08 | 0.21 | < 0.05 | 0.1556 | 0.09 | 0.16345 | 0.14 | 0.21 | 0.11 | [5 empty] | 7
| 1 | 6 | 1.1 | < 0.2 | 1.5 | < 0.1 | 0.08 | 0.23 | 0.16 | 0.0553 | 0.11 | 0.14903 | 0.13 | 0.23 | 0.12
| Mean |  | 1.1833 |  | 1.4667 |  | 0.081667 | 0.22 | 0.14667 | 0.09385 | 0.105 | 0.16307 | 0.135 | 0.22667 | 0.105
| Median |  | 1.2 |  | 1.45 |  | 0.08 | 0.21 | 0.16 | 0.0805 | 0.105 | 0.1611 | 0.135 | 0.23 | 0.11
| Std Dev. |  | 0.075277 |  | 0.08165 |  | 0.0040825 | 0.016733 | 0.032146 | 0.058644 | 0.010488 | 0.010153 | 0.010488 | 0.010328 | 0.016432
| Rel.Std.Dev. |  | 6.361% |  | 5.567% |  | 4.999% | 7.606% | 21.92% | 62.49% | 9.989% | 6.226% | 7.769% | 4.556% | 15.65%
| PDM3 |  | 738% |  | 938.6% |  | -42.17% | 55.79% | 3.862% | -33.54% | -25.64% | 15.48% | -4.4% | 60.51% | -25.64%
| Z-Score (Absolute) |  | 24.24 | 0.11 | 30.61 | 1.24 | 0.52 | 2.58 | 0.43 | 0.25 | 0 | 1.31 | 0.67 | 2.73 | 0
| Analytical results for Hf in OREAS 250c (Certified Value 4.34 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 4.39 | 5.1459 | 4.1517 | 4.1 | 0.1 | 3.7 | 4.3 | 4.2 | 4.9149 | 4.3 | 4.2419 | 4.31 | 4.3 | 4.5 | 4.3 | 4.19 |  |  | 1
| 1 | 2 | 4.59 | 5.4059 | 4.495 | 4.3 | 0.4 | 3.9 | 4.4 | 4.3 | 4.9492 | 4.3 | 4.2063 | 4.22 | 4.4 | 4.7 | 4.3 | 4 |  |  | 8
| 1 | 3 | 4.88 | 4.9779 | 4.474 | 4.3 | 0.6 | 4 | 4.6 | 4.2 | 5.0419 | 4.4 | 4.039 | 4.18 | 4.4 | 4.7 | 4.3 | 4.02 |  |  | 16
| 1 | 4 | 4.64 | 5.4461 | 4.3321 | 4.3 | 0.2 | 3.8 | 4.4 | 4.4 | 5.1294 | 4.1 | 3.9021 | 4.47 | 4.5 | 4.7 | 4.4 | 4.13 |  |  | 4.3446
| 1 | 5 | 4.46 | 5.1075 | 4.2852 | 4.2 | 0.1 | 4 | 4.5 | 4.4 | 4.7848 | 4.2 | 4.3626 | 4.11 | 4.5 | 4.5 | 4.5 | 3.91 |  |  | 34
| 1 | 6 | 4.27 | 5.2817 | 4.2616 | 4.4 | < 0.1 | 3.8 | 4.4 | 4.3 | 4.9855 | 4.4 | 3.7742 | 4.19 | 4.6 | 4.8 | 4.5 | 4.12
| Mean |  | 4.5383 | 5.2275 | 4.3333 | 4.2667 | 0.28 | 3.8667 | 4.4333 | 4.3 | 4.9676 | 4.2833 | 4.0877 | 4.2467 | 4.45 | 4.65 | 4.3833 | 4.0617
| Median |  | 4.525 | 5.2138 | 4.3086 | 4.3 | 0.2 | 3.85 | 4.4 | 4.3 | 4.9673 | 4.3 | 4.1227 | 4.205 | 4.45 | 4.7 | 4.35 | 4.07
| Std Dev. |  | 0.21442 | 0.18217 | 0.13142 | 0.10328 | 0.21679 | 0.12111 | 0.10328 | 0.089443 | 0.11709 | 0.1169 | 0.22295 | 0.12723 | 0.10488 | 0.12247 | 0.098319 | 0.10304
| Rel.Std.Dev. |  | 4.725% | 3.485% | 3.033% | 2.421% | 77.43% | 3.132% | 2.33% | 2.08% | 2.357% | 2.729% | 5.454% | 2.996% | 2.357% | 2.634% | 2.243% | 2.537%
| PDM3 |  | 4.46% | 20.32% | -0.2603% | -1.793% | -93.56% | -11% | 2.043% | -1.026% | 14.34% | -1.409% | -5.913% | -2.253% | 2.427% | 7.03% | 0.8924% | -6.512%
| Z-Score (Absolute) |  | 0.84 | 3.46 | 0.06 | 0.19 | 15.48 | 1.71 | 0.44 | 0.06 | 2.47 | 0.13 | 0.87 | 0.27 | 0.51 | 1.27 | 0.25 | 0.97
| Analytical results for Hg in OREAS 250c (Indicative Value 0.03 ppm) | [19 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.10 | 1.14 | [15 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [15 empty] | 3
3
| 1 | 1 | < 1 | < 2 | 0.0288 | [15 empty] | 1
| 1 | 2 | < 1 | < 2 | 0.0332 | [15 empty] | 2
| 1 | 3 | < 1 | < 2 | 0.0296 | [15 empty] | 16
| 1 | 4 | < 1 | < 2 | 0.02875 | [15 empty] | 0.030057
| 1 | 5 | < 1 | < 2 | 0.0313 | [15 empty] | 8
| 1 | 6 | < 1 | < 2 | 0.02869
| Mean |  |  |  | 0.030057
| Median |  |  |  | 0.0292
| Std Dev. |  |  |  | 0.0018324
| Rel.Std.Dev. |  |  |  | 6.097%
| PDM3 |  |  |  | -1.11e-13%
| Z-Score (Absolute) |  | 0 | 0.72 | 0.67
| Analytical results for Ho in OREAS 250c (Certified Value 0.63 ppm) | [19 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.61 | 0.65533 | 0.64059 | 0.6 | 0.62 | 0.6337 | 0.58 | [11 empty] | 1
| 1 | 2 | 0.65 | 0.66208 | 0.67908 | 0.7 | 0.61 | 0.6303 | 0.55 | [11 empty] | 4
| 1 | 3 | 0.65 | 0.62709 | 0.68239 | 0.7 | 0.62 | 0.6598 | 0.59 | [11 empty] | 16
| 1 | 4 | 0.63 | 0.67664 | 0.66753 | 0.6 | 0.6 | 0.6676 | 0.59 | [11 empty] | 0.62908
| 1 | 5 | 0.62 | 0.6353 | 0.67455 | 0.6 | 0.62 | 0.6284 | 0.56 | [11 empty] | 35
| 1 | 6 | 0.59 | 0.67237 | 0.62722 | 0.6 | 0.62 | 0.6468 | 0.57
| Mean |  | 0.625 | 0.6548 | 0.66189 | 0.63333 | 0.615 | 0.64443 | 0.57333
| Median |  | 0.625 | 0.65871 | 0.67104 | 0.6 | 0.62 | 0.64025 | 0.575
| Std Dev. |  | 0.023452 | 0.019935 | 0.022645 | 0.05164 | 0.0083666 | 0.016434 | 0.01633
| Rel.Std.Dev. |  | 3.752% | 3.044% | 3.421% | 8.154% | 1.36% | 2.55% | 2.848%
| PDM3 |  | -0.6481% | 4.089% | 5.217% | 0.6766% | -2.238% | 2.441% | -8.861%
| Z-Score (Absolute) |  | 0.33 | 0.68 | 0.92 | NA | 0.67 | 0.33 | 2.08
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 250c (Certified Value 0.059 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.06 | 0.06 | < 0.1 | < 2 | 0.055 | 0.06 | 0.0679 | 0.064 | 0.056039 | < 0.1 | 0.06 | 0.058 | 0.057 | 0.06 |  |  |  |  | 1
| 1 | 2 | 0.07 | 0.06 | < 0.1 | < 2 | 0.053 | 0.06 | 0.0667 | 0.07 | < 0.05 | < 0.1 | 0.064 | 0.061 | 0.047 | 0.057 |  |  |  |  | 22
| 1 | 3 | 0.07 | 0.06 | < 0.1 | < 2 | 0.07 | < 0.05 | 0.0724 | 0.068 | 0.05729 | < 0.1 | 0.059 | 0.062 | 0.06 | 0.065 |  |  |  |  | 16
| 1 | 4 | 0.07 | 0.06 | < 0.1 | < 2 | 0.057 | < 0.05 | 0.073 | 0.064 | 0.060448 | < 0.1 | 0.066 | 0.059 | 0.064 | 0.064 |  |  |  |  | 0.058553
| 1 | 5 | 0.06 | 0.05 | < 0.1 | < 2 | 0.06 | 0.05 | 0.0705 | 0.061 | < 0.05 | < 0.1 | 0.06 | 0.058 | 0.054 | 0.063 |  |  |  |  | 36
| 1 | 6 | 0.07 | 0.06 | < 0.1 | < 2 | 0.059 | 0.05 | 0.0694 | 0.067 | < 0.05 | < 0.1 | 0.057 | 0.059 | 0.058 | 0.057
| Mean |  | 0.066667 | 0.058333 |  |  | 0.059 | 0.055 | 0.069983 | 0.065667 | 0.057926 |  | 0.061 | 0.0595 | 0.056667 | 0.061
| Median |  | 0.07 | 0.06 |  |  | 0.058 | 0.055 | 0.06995 | 0.0655 | 0.05729 |  | 0.06 | 0.059 | 0.0575 | 0.0615
| Std Dev. |  | 0.005164 | 0.0040825 |  |  | 0.0059666 | 0.0057735 | 0.0024766 | 0.003266 | 0.0022727 |  | 0.0033466 | 0.0016432 | 0.005785 | 0.0035214
| Rel.Std.Dev. |  | 7.746% | 6.999% |  |  | 10.11% | 10.5% | 3.539% | 4.974% | 3.923% |  | 5.486% | 2.762% | 10.21% | 5.773%
| PDM3 |  | 13.86% | -0.3755% |  |  | 0.7631% | -6.068% | 19.52% | 12.15% | -1.072% |  | 4.179% | 1.617% | -3.222% | 4.179%
| Z-Score (Absolute) |  | 0.72 | 0.09 | 0.9 | 91.71 | 0.02 | 1.39 | 1.05 | 0.63 | 1.73 | 0.9 | 0.17 | 0.02 | 0.25 | 0.17
| Analytical results for K in OREAS 250c (Certified Value 1.32 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.3058 | 1.3605 | 1.396 | 1.2453 | 1.29 | 1.27 | 1.36 | 1.33 | 1.24 | 1.29 | 1.3548 | 1.66 | 1.3033 | 1.33 | 1.38 | 1.29 | 1.28 | 1.37 | 1
| 1 | 2 | 1.2766 | 1.4145 | 1.3809 | 1.2487 | 1.29 | 1.29 | 1.36 | 1.3 | 1.24 | 1.3 | 1.3237 | 1.29 | 1.3209 | 1.33 | 1.36 | 1.3 | 1.29 | 1.36 | #N/A
| 1 | 3 | 1.2735 | 1.3985 | 1.4052 | 1.2255 | 1.29 | 1.34 | 1.37 | 1.31 | 1.3 | 1.28 | 1.3247 | 1.32 | 1.3099 | 1.33 | 1.37 | 1.25 | 1.25 | 1.37 | 16
| 1 | 4 | 1.2953 | 1.4027 | 1.3605 | 1.2293 | 1.29 | 1.23 | 1.35 | 1.33 | 1.27 | 1.24 | 1.3607 | 1.3 | 1.3357 | 1.32 | 1.41 | 1.3 | 1.25 | 1.37 | 1.3185
| 1 | 5 | 1.308 | 1.4106 | 1.3609 | 1.2507 | 1.27 | 1.24 | 1.37 | 1.32 | 1.26 | 1.33 | 1.2839 | 1.42 | 1.2944 | 1.31 | 1.4 | 1.3 | 1.32 | 1.32 | 37
| 1 | 6 | 1.3086 | 1.435 | 1.3749 | 1.2549 | 1.29 | 1.18 | 1.37 | 1.34 | 1.26 | 1.25 | 1.3616 | 1.33 | 1.3468 | 1.34 | 1.42 | 1.33 | 1.28 | 1.34
| Mean |  | 1.2946 | 1.4036 | 1.3797 | 1.2424 | 1.2867 | 1.2583 | 1.3633 | 1.3217 | 1.2617 | 1.2817 | 1.3349 | 1.3867 | 1.3185 | 1.3267 | 1.39 | 1.295 | 1.2783 | 1.355
| Median |  | 1.3005 | 1.4066 | 1.3779 | 1.247 | 1.29 | 1.255 | 1.365 | 1.325 | 1.26 | 1.285 | 1.3398 | 1.325 | 1.3154 | 1.33 | 1.39 | 1.3 | 1.28 | 1.365
| Std Dev. |  | 0.015942 | 0.024656 | 0.018261 | 0.012076 | 0.008165 | 0.054924 | 0.008165 | 0.01472 | 0.022286 | 0.033116 | 0.030346 | 0.14166 | 0.019926 | 0.010328 | 0.023664 | 0.025884 | 0.026394 | 0.020736
| Rel.Std.Dev. |  | 1.231% | 1.757% | 1.323% | 0.972% | 0.6346% | 4.365% | 0.5989% | 1.114% | 1.766% | 2.584% | 2.273% | 10.22% | 1.511% | 0.7785% | 1.702% | 1.999% | 2.065% | 1.53%
| PDM3 |  | -1.811% | 6.455% | 4.643% | -5.772% | -2.415% | -4.564% | 3.4% | 0.2396% | -4.311% | -2.794% | 1.244% | 5.169% | 1.33e-05% | 0.6188% | 5.422% | -1.783% | -3.047% | 2.768%
| Z-Score (Absolute) |  | 0.43 | 1.41 | 1 | 1.31 | 0.56 | 1.04 | 0.73 | 0.03 | 0.98 | 0.65 | 0.25 | 1.12 | 0.03 | 0.11 | 1.18 | 0.42 | 0.7 | 0.59
| Analytical results for La in OREAS 250c (Certified Value 38 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 35.59 | 39.1541 | 39.3612 | 35.1 | 36 | 39.0372 | 37.2 | 36.4 | 32.2874 | 41.1 | 39.8 | 37.4 | 39.4 | 38.9 | 36.91 |  |  |  | 1
| 1 | 2 | 37.91 | 39.7801 | 39.6344 | 38 | 34 | 36.1512 | 36.8 | 37.3 | 31.4224 | 38.5 | 39 | 39.1 | 39.7 | 37.7 | 36.02 |  |  |  | 23
| 1 | 3 | 37.63 | 39.1612 | 39.9087 | 38.5 | 35 | 35.84 | 39.1 | 36.2 | 31.4228 | 40.4 | 38.8 | 38.3 | 39.4 | 38 | 36.46 |  |  |  | 16
| 1 | 4 | 38.06 | 40.8232 | 39.3232 | 35.7 | 34 | 37.9673 | 37.6 | 36.9 | 32.6361 | 36.7 | 39.9 | 40.7 | 39.7 | 39.3 | 36.44 |  |  |  | 37.9907
| 1 | 5 | 36.37 | 39.7021 | 38.4616 | 35.2 | 36 | 36.3454 | 38.4 | 37.4 | 30.4604 | 38.3 | 38.5 | 39.8 | 38 | 39.6 | 35.39 |  |  |  | 38
| 1 | 6 | 35.56 | 40.5175 | 39.5512 | 34.8 | 36 | 38.2394 | 37.6 | 37.2 | 32.136 | 39.6 | 39.3 | 40.7 | 39.5 | 40.2 | 36.68
| Mean |  | 36.8533 | 39.8564 | 39.3734 | 36.2167 | 35.1667 | 37.2634 | 37.7833 | 36.9 | 31.7275 | 39.1 | 39.2167 | 39.3333 | 39.2833 | 38.95 | 36.3167
| Median |  | 37 | 39.7411 | 39.4562 | 35.45 | 35.5 | 37.1563 | 37.6 | 37.05 | 31.7794 | 39.05 | 39.15 | 39.45 | 39.45 | 39.1 | 36.45
| Std Dev. |  | 1.1557 | 0.68961 | 0.49403 | 1.6092 | 0.98319 | 1.3191 | 0.83526 | 0.498 | 0.78707 | 1.5937 | 0.55648 | 1.3277 | 0.64317 | 0.95656 | 0.54158
| Rel.Std.Dev. |  | 3.136% | 1.73% | 1.255% | 4.443% | 2.796% | 3.54% | 2.211% | 1.35% | 2.481% | 4.076% | 1.419% | 3.375% | 1.637% | 2.456% | 1.491%
| PDM3 |  | -2.994% | 4.911% | 3.64% | -4.67% | -7.433% | -1.914% | -0.5458% | -2.871% | -16.49% | 2.92% | 3.227% | 3.534% | 3.402% | 2.525% | -4.406%
| Z-Score (Absolute) |  | 0.43 | 0.95 | 0.73 | 0.72 | 1.2 | 0.24 | 0 | 0.41 | 2.78 | 0.61 | 0.66 | 0.71 | 0.69 | 0.54 | 0.67
| Analytical results for Li in OREAS 250c (Certified Value 21.7 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 20.7 | 23.1909 | 22.7633 | 35.9062 | 19.9 | 22 | 22.1 | 20.402 | 20.4 | 21.4 | 18.876 | 22.5 | 21.7388 | 21.5 | 21.9 | 22.1 | 21.8 | 22 | 1
| 1 | 2 | 21.5 | 23.6912 | 22.805 | 34.9667 | 20.5 | 23 | 22.1 | 19.5312 | 20.3 | 20.6 | 18.174 | 21.7 | 21.1674 | 20.9 | 21.9 | 22.5 | 22.2 | 22 | #N/A
| 1 | 3 | 22.4 | 23.5328 | 22.6278 | 35.8799 | 20.2 | 22 | 21.9 | 19.4756 | 21.5 | 21.8 | 18.408 | 22.1 | 21.6692 | 20.8 | 21.9 | 21.6 | 21.5 | 23 | 16
| 1 | 4 | 21.5 | 24.589 | 22.2314 | 35.6721 | 20.6 | 23 | 22 | 20.7136 | 20.9 | 21.5 | 19.032 | 22.1 | 21.1827 | 20.7 | 22.6 | 22.4 | 21.4 | 23 | 21.7332
| 1 | 5 | 20.5 | 23.8021 | 22.4944 | 36.4957 | 20.5 | 23 | 22.5 | 19.5171 | 20.6 | 20.5 | 17.706 | 22.6 | 20.8231 | 19.9 | 22.6 | 22.9 | 22.8 | 22 | 39
| 1 | 6 | 20 | 24.1607 | 22.475 | 36.4553 | 20.5 | 23 | 22.6 | 20.7862 | 20.9 | 20.6 | 18.798 | 22.5 | 21.3207 | 20.4 | 22.7 | 22.5 | 21.7 | 22
| Mean |  | 21.1 | 23.8278 | 22.5661 | 35.896 | 20.3667 | 22.6667 | 22.2 | 20.0709 | 20.7667 | 21.0667 | 18.499 | 22.25 | 21.317 | 20.7 | 22.2667 | 22.3333 | 21.9 | 22.3333
| Median |  | 21.1 | 23.7466 | 22.5611 | 35.8931 | 20.5 | 23 | 22.1 | 19.9666 | 20.75 | 21 | 18.603 | 22.3 | 21.2517 | 20.75 | 22.25 | 22.45 | 21.75 | 22
| Std Dev. |  | 0.86487 | 0.49036 | 0.21225 | 0.5633 | 0.26583 | 0.5164 | 0.28284 | 0.63035 | 0.43665 | 0.56451 | 0.50167 | 0.34496 | 0.34242 | 0.53292 | 0.40332 | 0.44121 | 0.52154 | 0.5164
| Rel.Std.Dev. |  | 4.099% | 2.058% | 0.9406% | 1.569% | 1.305% | 2.278% | 1.274% | 3.141% | 2.103% | 2.68% | 2.712% | 1.55% | 1.606% | 2.574% | 1.811% | 1.976% | 2.381% | 2.312%
| PDM3 |  | -2.914% | 9.637% | 3.832% | 65.17% | -6.288% | 4.295% | 2.148% | -7.649% | -4.447% | -3.067% | -14.88% | 2.378% | -1.915% | -4.754% | 2.454% | 2.761% | 0.7673% | 2.761%
| Z-Score (Absolute) |  | 0.76 | 1.42 | 0.41 | 11.09 | 1.35 | 0.49 | 0.12 | 1.59 | 1.03 | 0.79 | 2.85 | 0.16 | 0.59 | 1.08 | 0.17 | 0.23 | 0.12 | 0.23
| Analytical results for Lu in OREAS 250c (Certified Value 0.18 ppm) | [19 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.17 | 0.18997 | 0.18844 | 0.2 | 0.18 | 0.0971 | 0.14 | [11 empty] | 1
| 1 | 2 | 0.19 | 0.19457 | 0.19203 | 0.2 | 0.18 | 0.0965 | 0.14 | [11 empty] | 5
| 1 | 3 | 0.19 | 0.19049 | 0.19538 | 0.2 | 0.18 | 0.0871 | 0.14 | [11 empty] | 16
| 1 | 4 | 0.19 | 0.19644 | 0.21308 | 0.2 | 0.18 | 0.0989 | 0.14 | [11 empty] | 0.17769
| 1 | 5 | 0.18 | 0.19338 | 0.19277 | 0.2 | 0.18 | 0.0912 | 0.13 | [11 empty] | 40
| 1 | 6 | 0.17 | 0.20685 | 0.17741 | 0.2 | 0.19 | 0.0968 | 0.13
| Mean |  | 0.18167 | 0.19528 | 0.19319 | 0.2 | 0.18167 | 0.0946 | 0.13667
| Median |  | 0.185 | 0.19397 | 0.1924 | 0.2 | 0.18 | 0.09665 | 0.14
| Std Dev. |  | 0.0098319 | 0.0061742 | 0.011605 | 3.0405e-17 | 0.0040825 | 0.0044944 | 0.005164
| Rel.Std.Dev. |  | 5.412% | 3.162% | 6.007% | 1.52e-14% | 2.247% | 4.751% | 3.779%
| PDM3 |  | 2.236% | 9.898% | 8.719% | 12.55% | 2.236% | -46.76% | -23.09%
| Z-Score (Absolute) |  | 0 | 0.73 | 0.62 | NA | 0 | 4.67 | 2.41
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 250c (Certified Value 1.55 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.5717 | 1.6699 | 1.6256 | 1.7417 | 1.57 | 1.47 | 1.52 | 1.56 | 1.48 | 1.53 | 1.6057 | 1.6 | 1.5669 | 1.62 | 1.53 | 1.53 | 1.49 | 1.477 | 1
| 1 | 2 | 1.5335 | 1.7436 | 1.622 | 1.7533 | 1.56 | 1.66 | 1.53 | 1.52 | 1.48 | 1.58 | 1.5462 | 1.55 | 1.5845 | 1.61 | 1.54 | 1.56 | 1.5 | 1.497 | #N/A
| 1 | 3 | 1.5304 | 1.7115 | 1.6305 | 1.7003 | 1.58 | 1.67 | 1.54 | 1.51 | 1.56 | 1.56 | 1.5673 | 1.57 | 1.541 | 1.6 | 1.55 | 1.51 | 1.47 | 1.489 | 16
| 1 | 4 | 1.5676 | 1.6973 | 1.6172 | 1.7045 | 1.57 | 1.51 | 1.53 | 1.58 | 1.54 | 1.59 | 1.6164 | 1.55 | 1.5774 | 1.6 | 1.58 | 1.56 | 1.47 | 1.52 | 1.5542
| 1 | 5 | 1.5915 | 1.7643 | 1.6096 | 1.7488 | 1.58 | 1.42 | 1.55 | 1.55 | 1.51 | 1.62 | 1.5222 | 1.61 | 1.5455 | 1.58 | 1.57 | 1.54 | 1.56 | 1.479 | 41
| 1 | 6 | 1.4824 | 1.7749 | 1.6264 | 1.7438 | 1.58 | 1.41 | 1.52 | 1.58 | 1.53 | 1.6 | 1.6122 | 1.6 | 1.5781 | 1.62 | 1.58 | 1.58 | 1.51 | 1.478
| Mean |  | 1.5462 | 1.7269 | 1.6219 | 1.7321 | 1.5733 | 1.5233 | 1.5317 | 1.55 | 1.5167 | 1.58 | 1.5783 | 1.58 | 1.5656 | 1.605 | 1.5583 | 1.5467 | 1.5 | 1.49
| Median |  | 1.5505 | 1.7275 | 1.6238 | 1.7427 | 1.575 | 1.49 | 1.53 | 1.555 | 1.52 | 1.585 | 1.5865 | 1.585 | 1.5722 | 1.605 | 1.56 | 1.55 | 1.495 | 1.484
| Std Dev. |  | 0.039094 | 0.040839 | 0.0074982 | 0.023381 | 0.008165 | 0.11553 | 0.01169 | 0.029665 | 0.03266 | 0.031623 | 0.039112 | 0.026833 | 0.018245 | 0.015166 | 0.02137 | 0.025033 | 0.033466 | 0.016637
| Rel.Std.Dev. |  | 2.528% | 2.365% | 0.4623% | 1.35% | 0.519% | 7.584% | 0.7633% | 1.914% | 2.153% | 2.001% | 2.478% | 1.698% | 1.165% | 0.9449% | 1.371% | 1.619% | 2.231% | 1.117%
| PDM3 |  | -0.5149% | 11.11% | 4.355% | 11.44% | 1.232% | -1.985% | -1.449% | -0.2693% | -2.414% | 1.661% | 1.554% | 1.661% | 0.7326% | 3.27% | 0.2669% | -0.4838% | -3.486% | -4.13%
| Z-Score (Absolute) |  | 0.44 | 4.6 | 1.67 | 4.75 | 0.32 | 1.08 | 0.85 | 0.33 | 1.26 | 0.5 | 0.46 | 0.5 | 0.1 | 1.2 | 0.1 | 0.43 | 1.73 | 2.01
| Analytical results for Mn in OREAS 250c (Certified Value 0.049 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0515 | 0.048443 | 0.051118 | < 0.1 | 0.048 | 0.0445 | 0.0487 | 0.0492 | 0.0472 | 0.0474 | 0.04454 | 0.0509 | 0.050924 | 0.0508 | 0.0489 | 0.0504 | 0.0487 | 0.0468 | 1
| 1 | 2 | 0.0502 | 0.048925 | 0.051213 | < 0.1 | 0.048 | 0.0486 | 0.0503 | 0.0485 | 0.0469 | 0.0487 | 0.043357 | 0.049 | 0.048143 | 0.0501 | 0.0494 | 0.0509 | 0.0485 | 0.0469 | 25
| 1 | 3 | 0.0499 | 0.04841 | 0.050938 | < 0.1 | 0.0485 | 0.0495 | 0.0496 | 0.0503 | 0.0493 | 0.048 | 0.043572 | 0.0506 | 0.049288 | 0.0502 | 0.0494 | 0.0488 | 0.0474 | 0.0466 | 16
| 1 | 4 | 0.0512 | 0.048348 | 0.050884 | < 0.1 | 0.0483 | 0.0457 | 0.049 | 0.0499 | 0.0483 | 0.0488 | 0.044756 | 0.0489 | 0.049596 | 0.0501 | 0.0507 | 0.0508 | 0.0487 | 0.0476 | 0.049236
| 1 | 5 | 0.0522 | 0.04842 | 0.050829 | < 0.1 | 0.0485 | 0.0441 | 0.0498 | 0.049 | 0.0477 | 0.0495 | 0.042605 | 0.0511 | 0.052038 | 0.0485 | 0.0496 | 0.0499 | 0.0508 | 0.0466 | 42
| 1 | 6 | 0.0485 | 0.04808 | 0.051159 | < 0.1 | 0.0491 | 0.05 | 0.0502 | 0.0497 | 0.0482 | 0.049 | 0.044972 | 0.0502 | 0.052389 | 0.0498 | 0.0506 | 0.0516 | 0.0491 | 0.0462
| Mean |  | 0.050583 | 0.048438 | 0.051023 |  | 0.0484 | 0.047067 | 0.0496 | 0.049433 | 0.047933 | 0.048567 | 0.043967 | 0.050117 | 0.050396 | 0.049917 | 0.049767 | 0.0504 | 0.048867 | 0.046783
| Median |  | 0.0507 | 0.048415 | 0.051028 |  | 0.0484 | 0.04715 | 0.0497 | 0.04945 | 0.04795 | 0.04875 | 0.044056 | 0.0504 | 0.05026 | 0.0501 | 0.0495 | 0.0506 | 0.0487 | 0.0467
| Std Dev. |  | 0.0013258 | 0.00027382 | 0.00016005 |  | 0.00040988 | 0.0026128 | 0.00064187 | 0.00065625 | 0.0008641 | 0.00075011 | 0.00093203 | 0.00095376 | 0.0016666 | 0.0007679 | 0.00072296 | 0.0009654 | 0.0011075 | 0.00046655
| Rel.Std.Dev. |  | 2.621% | 0.5653% | 0.3137% |  | 0.8469% | 5.551% | 1.294% | 1.328% | 1.803% | 1.544% | 2.12% | 1.903% | 3.307% | 1.538% | 1.453% | 1.915% | 2.266% | 0.9973%
| PDM3 |  | 2.736% | -1.622% | 3.63% |  | -1.699% | -4.407% | 0.7387% | 0.4002% | -2.646% | -1.36% | -10.7% | 1.788% | 2.356% | 1.382% | 1.077% | 2.364% | -0.7507% | -4.982%
| Z-Score (Absolute) |  | 0.78 | 0.79 | 1.11 | 0.36 | 0.82 | 1.8 | 0.06 | 0.06 | 1.16 | 0.7 | 4.08 | 0.44 | 0.65 | 0.29 | 0.18 | 0.65 | 0.48 | 2.01
| Analytical results for Mo in OREAS 250c (Certified Value 1.96 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2 | 1.9854 | 3.628 | 2 | < 1 | 2.09 | 2.065 | 1.93 | 1.8 | 2.1383 | 1.9 | 1.8093 | 1.8 | 1.83 | 2 | 1.96 | 2.24 |  | 1
| 1 | 2 | 1.9 | 1.9566 | 3.6511 | 2 | < 1 | 2.07 | 2.3287 | 1.93 | 1.8 | 2.1106 | 1.8 | 1.735 | 1.73 | 1.86 | 2.06 | 1.94 | 2.37 |  | 26
| 1 | 3 | 1.9 | 2.0123 | 3.468 | 2 | < 1 | 2.09 | 2.0602 | 1.98 | 1.8 | 2.1417 | 1.98 | 1.5493 | 1.75 | 1.91 | 2.04 | 1.94 | 2.38 |  | 16
| 1 | 4 | 2 | 1.9897 | 3.4641 | 2 | < 1 | 2.07 | < 2 | 2 | 1.8 | 2.2239 | 1.65 | 1.6122 | 1.76 | 1.9 | 2.02 | 1.96 | 2.23 |  | 1.9642
| 1 | 5 | 1.9 | 1.9138 | 3.8084 | 2 | < 1 | 2.09 | < 2 | 2.05 | 1.9 | 2.1539 | 1.79 | 1.7044 | 1.72 | 2 | 2 | 2.04 | 2.22 |  | 43
| 1 | 6 | 1.8 | 2.0577 | 3.7841 | 1.9 | < 1 | 2.08 | < 2 | 1.84 | 1.8 | 2.1613 | 1.92 | 1.5624 | 1.77 | 1.9 | 2.07 | 2.11 | 2.21
| Mean |  | 1.9167 | 1.9859 | 3.634 | 1.9833 |  | 2.0817 | 2.1513 | 1.955 | 1.8167 | 2.1549 | 1.84 | 1.6621 | 1.755 | 1.9 | 2.0317 | 1.9917 | 2.275
| Median |  | 1.9 | 1.9875 | 3.6395 | 2 |  | 2.085 | 2.065 | 1.955 | 1.8 | 2.1478 | 1.85 | 1.6583 | 1.755 | 1.9 | 2.03 | 1.96 | 2.235
| Std Dev. |  | 0.075277 | 0.048855 | 0.14812 | 0.040825 |  | 0.0098319 | 0.15364 | 0.072319 | 0.040825 | 0.037987 | 0.11815 | 0.10386 | 0.02881 | 0.057619 | 0.029944 | 0.068823 | 0.078166
| Rel.Std.Dev. |  | 3.928% | 2.46% | 4.076% | 2.058% |  | 0.4723% | 7.142% | 3.699% | 2.247% | 1.763% | 6.421% | 6.249% | 1.642% | 3.033% | 1.474% | 3.456% | 3.436%
| PDM3 |  | -2.422% | 1.105% | 85.01% | 0.9724% |  | 5.979% | 9.523% | -0.4701% | -7.513% | 9.709% | -6.325% | -15.38% | -10.65% | -3.27% | 3.433% | 1.397% | 15.82%
| Z-Score (Absolute) |  | 0.2 | 0.16 | 8.94 | 0.15 | 7.75 | 0.67 | 2.02 | 0 | 0.74 | 1.06 | 0.61 | 1.56 | 1.06 | 0.29 | 0.41 | 0.2 | 1.7
| Analytical results for Na in OREAS 250c (Certified Value 1.07 wt.%) | [19 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.0832 | 0.96579 | 1.1336 | 2.6176 | 1.06 | 0.98 | 1.07 | 1.11 | 1.02 | 1.09 | 1.1158 | 1.09 | 1.0757 | 1.09 | 1.05 | 1.06 | 1.04 | 1.07 | 1
| 1 | 2 | 1.0659 | 0.9616 | 1.1332 | 2.5959 | 1.07 | 1.11 | 1.07 | 1.08 | 1.01 | 1.086 | 1.0799 | 1.05 | 1.0723 | 1.09 | 1.05 | 1.08 | 1.05 | 1.062 | #N/A
| 1 | 3 | 1.0678 | 0.95268 | 1.1265 | 2.5755 | 1.07 | 1.1 | 1.11 | 1.09 | 1.06 | 1.064 | 1.0838 | 1.07 | 1.0914 | 1.08 | 1.06 | 1.05 | 1.03 | 1.077 | 16
| 1 | 4 | 1.0908 | 0.95666 | 1.1036 | 2.5628 | 1.06 | 1.03 | 1.1 | 1.1 | 1.04 | 1.076 | 1.1213 | 1.05 | 1.0814 | 1.09 | 1.08 | 1.08 | 1.03 | 1.075 | 1.0741
| 1 | 5 | 1.0985 | 0.96765 | 1.1189 | 2.6015 | 1.07 | 0.99 | 1.08 | 1.08 | 1.04 | 1.104 | 1.0485 | 1.09 | 1.0865 | 1.07 | 1.08 | 1.07 | 1.09 | 1.045 | 44
| 1 | 6 | 1.0309 | 0.96197 | 1.1232 | 2.6076 | 1.07 | 0.95 | 1.08 | 1.13 | 1.04 | 1.082 | 1.1237 | 1.1 | 1.0888 | 1.1 | 1.09 | 1.1 | 1.06 | 1.054
| Mean |  | 1.0729 | 0.96106 | 1.1232 | 2.5935 | 1.0667 | 1.0267 | 1.085 | 1.0983 | 1.035 | 1.0837 | 1.0955 | 1.075 | 1.0827 | 1.0867 | 1.0683 | 1.0733 | 1.05 | 1.0638
| Median |  | 1.0755 | 0.96178 | 1.1248 | 2.5987 | 1.07 | 1.01 | 1.08 | 1.095 | 1.04 | 1.084 | 1.0998 | 1.08 | 1.0839 | 1.09 | 1.07 | 1.075 | 1.045 | 1.066
| Std Dev. |  | 0.02417 | 0.0055944 | 0.011165 | 0.020556 | 0.005164 | 0.065929 | 0.016432 | 0.019408 | 0.017607 | 0.013471 | 0.029874 | 0.021679 | 0.007549 | 0.010328 | 0.017224 | 0.017512 | 0.022804 | 0.012576
| Rel.Std.Dev. |  | 2.253% | 0.5821% | 0.994% | 0.7926% | 0.4841% | 6.422% | 1.514% | 1.767% | 1.701% | 1.243% | 2.727% | 2.017% | 0.6973% | 0.9504% | 1.612% | 1.632% | 2.172% | 1.182%
| PDM3 |  | -0.1128% | -10.52% | 4.573% | 141.5% | -0.6885% | -4.413% | 1.018% | 2.26% | -3.637% | 0.8943% | 1.996% | 0.08739% | 0.8021% | 1.174% | -0.5333% | -0.06779% | -2.24% | -0.9523%
| Z-Score (Absolute) |  | 0.08 | 6.54 | 2.83 | 87.81 | 0.43 | 2.75 | 0.63 | 1.4 | 2.26 | 0.55 | 1.23 | 0.05 | 0.49 | 0.72 | 0.34 | 0.05 | 1.4 | 0.6
| Analytical results for Nb in OREAS 250c (Certified Value 31.8 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 31.37 | 31.8196 | 85.1626 | 31 | 0.8 | 28 | 28.546 | 30.1 | 32.1 | 49.0084 | 33.2 | 35.8917 | 31.2 | 33.3 | 32.8 | 31.6 | 33.8 |  | 1
| 1 | 2 | 32.23 | 32.8018 | 53.8107 | 30.9 | 0.2 | 26 | 28.5413 | 29.3 | 33 | 48.9774 | 32.3 | 34.5227 | 30.7 | 33.2 | 33.8 | 31.1 | 31.8 |  | 27
| 1 | 3 | 32.22 | 33.4553 | 60.6103 | 30.6 | 0.2 | 29 | 29.213 | 31.3 | 32.4 | 50.2144 | 33.4 | 33.7425 | 30.9 | 33 | 33.7 | 31.6 | 33.4 |  | 16
| 1 | 4 | 32.46 | 33.0263 | 58.0611 | 31.4 | 0.2 | 28 | 29.3554 | 30.3 | 32.2 | 51.4919 | 30.5 | 33.3825 | 31.4 | 33.8 | 33.6 | 31.2 | 32.4 |  | 31.7716
| 1 | 5 | 31.86 | 33.6542 | 63.9197 | 30.4 | 0.5 | 29 | 29.254 | 30.2 | 32.8 | 47.8759 | 31.4 | 35.6787 | 29.3 | 33.7 | 32.5 | 31.9 | 32.4 |  | 45
| 1 | 6 | 30.38 | 31.3389 | 61.209 | 30.9 | 0.3 | 29 | 29.9207 | 30 | 30.8 | 49.3834 | 32.5 | 35.2219 | 30.8 | 34.5 | 34.1 | 33.5 | 32
| Mean |  | 31.7533 | 32.6827 | 63.7956 | 30.8667 | 0.36667 | 28.1667 | 29.1384 | 30.2 | 32.2167 | 49.4919 | 32.2167 | 34.74 | 30.7167 | 33.5833 | 33.4167 | 31.8167 | 32.6333
| Median |  | 32.04 | 32.914 | 60.9096 | 30.9 | 0.25 | 28.5 | 29.2335 | 30.15 | 32.3 | 49.1959 | 32.4 | 34.8723 | 30.85 | 33.5 | 33.65 | 31.6 | 32.4
| Std Dev. |  | 0.77314 | 0.91907 | 11.0089 | 0.34448 | 0.24221 | 1.169 | 0.52642 | 0.64498 | 0.77567 | 1.2364 | 1.1017 | 1.032 | 0.7414 | 0.54191 | 0.62423 | 0.87502 | 0.79415
| Rel.Std.Dev. |  | 2.435% | 2.812% | 17.26% | 1.116% | 66.06% | 4.15% | 1.807% | 2.136% | 2.408% | 2.498% | 3.42% | 2.971% | 2.414% | 1.614% | 1.868% | 2.75% | 2.434%
| PDM3 |  | -0.0575% | 2.868% | 100.8% | -2.848% | -98.85% | -11.35% | -8.288% | -4.947% | 1.401% | 55.77% | 1.401% | 9.343% | -3.32% | 5.702% | 5.178% | 0.1418% | 2.712%
| Z-Score (Absolute) |  | 0.23 | 0.23 | 15.58 | 0.67 | 15.71 | 2 | 1.52 | 1 | 0 | 8.52 | 0 | 1.25 | 0.74 | 0.67 | 0.59 | 0.2 | 0.21
| Analytical results for Nd in OREAS 250c (Certified Value 32.1 ppm) | [19 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.09 | 1.10 | 1.14 | 1.22 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 1
2
| 1 | 1 | 31.08 | 34.7172 | 31.5 | 30.5 | 26.4564 | 32.2 | [12 empty] | 1
| 1 | 2 | 32.74 | 34.4273 | 34.1 | 30.7 | 25.6744 | 31.1 | [12 empty] | 6
| 1 | 3 | 32.01 | 34.8325 | 34.4 | 30.2 | 25.7915 | 31.6 | [12 empty] | 16
| 1 | 4 | 32.7 | 34.833 | 31.6 | 29.7 | 26.6699 | 31.8 | [12 empty] | 32.0964
| 1 | 5 | 31 | 34.6859 | 31.6 | 30.9 | 24.942 | 30.2 | [12 empty] | 46
| 1 | 6 | 30.55 | 35.0174 | 30.9 | 30.8 | 26.4393 | 30.5
| Mean |  | 31.68 | 34.7522 | 32.35 | 30.4667 | 25.9956 | 31.2333
| Median |  | 31.545 | 34.7749 | 31.6 | 30.6 | 26.1154 | 31.35
| Std Dev. |  | 0.93511 | 0.19721 | 1.4977 | 0.45019 | 0.65098 | 0.77632
| Rel.Std.Dev. |  | 2.952% | 0.5675% | 4.63% | 1.478% | 2.504% | 2.486%
| PDM3 |  | -1.297% | 8.274% | 0.79% | -5.078% | -19.01% | -2.689%
| Z-Score (Absolute) |  | 0.16 | 2.36 | 0.64 | 0.71 | 3.91 | 0.16
| Analytical results for Ni in OREAS 250c (Certified Value 83 ppm) | [19 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
1
| 1 | 1 | 82.6 | 88.9935 | 72.7007 | 80 | 87 | 83 | 80.755 | 80 | 79.1 | 69.4671 | 84.8 | 85.4652 | 84.5 | 83.7 | 85.9 | 83.8 | 89 |  | 1
| 1 | 2 | 84 | 87.9312 | 74.1995 | 81 | 89 | 82 | 78.03 | 78 | 81.7 | 67.8246 | 79.8 | 83.7569 | 83.8 | 84.1 | 88.3 | 82.7 | 87 |  | 29
| 1 | 3 | 85 | 87.6619 | 71.4333 | 81 | 86 | 83 | 78.205 | 84.1 | 80.4 | 68.4876 | 82.5 | 83.3317 | 83.4 | 82.8 | 87.2 | 82.7 | 90 |  | 16
| 1 | 4 | 85.8 | 87.8112 | 72.7041 | 80 | 88 | 83 | 80.1733 | 81.7 | 81.7 | 69.957 | 76.6 | 83.7606 | 86.8 | 86 | 87.8 | 83.4 | 88 |  | 83.1882
| 1 | 5 | 83.9 | 91.1179 | 73.1362 | 80 | 88 | 85 | 78.815 | 80.3 | 82.7 | 66.5251 | 79.9 | 87.0381 | 84.2 | 85.4 | 85.1 | 87.2 | 84 |  | 47
| 1 | 6 | 81.2 | 87.2462 | 74.073 | 80 | 88 | 83 | 80.2367 | 80.1 | 90.7 | 70.0767 | 82.5 | 86.2643 | 86.4 | 86.5 | 86.5 | 85.7 | 86
| Mean |  | 83.75 | 88.4603 | 73.0411 | 80.3333 | 87.6667 | 83.1667 | 79.3692 | 80.7 | 82.7167 | 68.723 | 81.0167 | 84.9361 | 84.85 | 84.75 | 86.8 | 84.25 | 87.3333
| Median |  | 83.95 | 87.8712 | 72.9202 | 80 | 88 | 83 | 79.4942 | 80.2 | 81.7 | 68.9774 | 81.2 | 84.6129 | 84.35 | 84.75 | 86.85 | 83.6 | 87.5
| Std Dev. |  | 1.6538 | 1.4255 | 1.0228 | 0.5164 | 1.0328 | 0.98319 | 1.1641 | 2.0425 | 4.1048 | 1.3858 | 2.8618 | 1.5368 | 1.411 | 1.4405 | 1.2 | 1.8185 | 2.1602
| Rel.Std.Dev. |  | 1.975% | 1.611% | 1.4% | 0.6428% | 1.178% | 1.182% | 1.467% | 2.531% | 4.963% | 2.017% | 3.532% | 1.809% | 1.663% | 1.7% | 1.382% | 2.158% | 2.474%
| PDM3 |  | 0.6753% | 6.337% | -12.2% | -3.432% | 5.383% | -0.02594% | -4.591% | -2.991% | -0.5669% | -17.39% | -2.61% | 2.101% | 1.998% | 1.877% | 4.342% | 1.276% | 4.983%
| Z-Score (Absolute) |  | 0 | 1.04 | 2.37 | 0.76 | 0.87 | 0.13 | 0.97 | 0.67 | 0.23 | 3.32 | 0.6 | 0.26 | 0.24 | 0.22 | 0.67 | 0.11 | 0.79
| Analytical results for P in OREAS 250c (Certified Value 0.095 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0944 | 0.099283 | < 0.1 | 0.099 | 0.081 | 0.0991 | 0.0938 | 0.091 | 0.093 | 0.090779 | 0.098 | 0.097168 | 0.1 | 0.094 | 0.095 | 0.091 | 0.092 |  | 1
| 1 | 2 | 0.0922 | 0.099028 | < 0.1 | 0.098 | 0.079 | 0.0997 | 0.0916 | 0.091 | 0.096 | 0.088083 | 0.095 | 0.098761 | 0.1 | 0.093 | 0.097 | 0.092 | 0.093 |  | #N/A
| 1 | 3 | 0.0923 | 0.098662 | < 0.1 | 0.097 | 0.077 | 0.0995 | 0.09 | 0.095 | 0.093 | 0.089082 | 0.098 | 0.096955 | 0.1 | 0.095 | 0.095 | 0.09 | 0.092 |  | 16
| 1 | 4 | 0.0938 | 0.09773 | < 0.1 | 0.098 | 0.078 | 0.0988 | 0.0937 | 0.094 | 0.095 | 0.091479 | 0.095 | 0.09962 | 0.1 | 0.098 | 0.096 | 0.09 | 0.094 |  | 0.095295
| 1 | 5 | 0.0955 | 0.097932 | < 0.1 | 0.097 | 0.08 | 0.0994 | 0.0929 | 0.092 | 0.098 | 0.086585 | 0.098 | 0.096021 | 0.1 | 0.097 | 0.096 | 0.095 | 0.091 |  | 48
| 1 | 6 | 0.0892 | 0.09733 | < 0.1 | 0.098 | 0.08 | 0.0996 | 0.0926 | 0.093 | 0.095 | 0.090579 | 0.098 | 0.10035 | 0.1 | 0.098 | 0.099 | 0.093 | 0.092
| Mean |  | 0.0929 | 0.098327 |  | 0.097833 | 0.079167 | 0.09935 | 0.092433 | 0.092667 | 0.095 | 0.089431 | 0.097 | 0.098146 | 0.1 | 0.095833 | 0.096333 | 0.091833 | 0.092333
| Median |  | 0.09305 | 0.098297 |  | 0.098 | 0.0795 | 0.09945 | 0.09275 | 0.0925 | 0.095 | 0.08983 | 0.098 | 0.097964 | 0.1 | 0.096 | 0.096 | 0.0915 | 0.092
| Std Dev. |  | 0.0022073 | 0.0007778 |  | 0.00075277 | 0.001472 | 0.00033912 | 0.0014376 | 0.001633 | 0.0018974 | 0.0018634 | 0.0015492 | 0.0016911 | 1.5202e-17 | 0.002137 | 0.0015055 | 0.0019408 | 0.0010328
| Rel.Std.Dev. |  | 2.376% | 0.791% |  | 0.7694% | 1.859% | 0.3413% | 1.555% | 1.762% | 1.997% | 2.084% | 1.597% | 1.723% | 1.52e-14% | 2.23% | 1.563% | 2.113% | 1.119%
| PDM3 |  | -2.513% | 3.182% |  | 2.664% | -16.92% | 4.256% | -3.003% | -2.758% | -0.3093% | -6.153% | 1.789% | 2.992% | 4.938% | 0.5652% | 1.09% | -3.632% | -3.108%
| Z-Score (Absolute) |  | 0.5 | 0.79 | 10.71 | 0.67 | 3.77 | 1.04 | 0.61 | 0.56 | 0 | 1.33 | 0.48 | 0.75 | 1.19 | 0.2 | 0.32 | 0.75 | 0.63
| Analytical results for Pb in OREAS 250c (Certified Value 9.92 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 10.8 | 11.3066 | 5.2321 | 11 | 9 | 10.1 | 8.6775 | 10 | 9.9 | 8.2575 | 10 | 9.9399 | 10.9 | 9.5 | 10.4 | 10.4 |  |  | 1
| 1 | 2 | 11.1 | 15.2722 | 6.0214 | 11 | 10 | 10.3 | 7.6278 | 9.9 | 10.1 | 5.5138 | 9.8 | 9.3554 | 10.8 | 9.6 | 10.7 | 10.5 |  |  | 30
| 1 | 3 | 10.8 | 10.6533 | 4.9826 | 11 | 11 | 9.8 | 7.5746 | 10.5 | 10.1 | 8.2964 | 10.2 | 9.7352 | 10.6 | 9.7 | 10.7 | 10.2 |  |  | 16
| 1 | 4 | 11.3 | 11.2827 | 8.164 | 11 | 10 | 10.2 | 8.4663 | 10 | 10 | 8.0343 | 9.4 | 8.972 | 10.7 | 10 | 10.8 | 10.4 |  |  | 9.9208
| 1 | 5 | 10.6 | 10.6793 | 8.0335 | 11 | 9 | 9.7 | 8.03 | 10 | 10 | 7.9989 | 9.7 | 10.3719 | 10.5 | 9.9 | 10.4 | 10.8 |  |  | 49
| 1 | 6 | 10.2 | 11.0822 | 5.3084 | 11 | 10 | 9.6 | 8.4286 | 10.3 | 9.7 | 7.9911 | 9.9 | 8.7016 | 10.6 | 10 | 10.7 | 10.9
| Mean |  | 10.8 | 11.7127 | 6.2903 | 11 | 9.8333 | 9.95 | 8.1341 | 10.1167 | 9.9667 | 7.682 | 9.8333 | 9.5127 | 10.6833 | 9.7833 | 10.6167 | 10.5333
| Median |  | 10.8 | 11.1825 | 5.6649 | 11 | 10 | 9.95 | 8.2293 | 10 | 10 | 8.0166 | 9.85 | 9.5453 | 10.65 | 9.8 | 10.7 | 10.45
| Std Dev. |  | 0.38471 | 1.7668 | 1.4433 | 0 | 0.75277 | 0.2881 | 0.46315 | 0.23166 | 0.15055 | 1.0705 | 0.27325 | 0.62374 | 0.1472 | 0.2137 | 0.17224 | 0.26583
| Rel.Std.Dev. |  | 3.562% | 15.08% | 22.94% | 0% | 7.655% | 2.895% | 5.694% | 2.29% | 1.511% | 13.94% | 2.779% | 6.557% | 1.378% | 2.184% | 1.622% | 2.524%
| PDM3 |  | 8.863% | 18.06% | -36.59% | 10.88% | -0.8813% | 0.2947% | -18.01% | 1.975% | 0.4627% | -22.57% | -0.8813% | -4.114% | 7.687% | -1.385% | 7.015% | 6.175%
| Z-Score (Absolute) |  | 0.92 | 1.92 | 4.01 | NA | NA | 0.01 | 1.99 | 0.17 | 0.01 | 2.49 | 0.14 | 0.49 | 0.79 | 0.19 | 0.72 | 0.63
| Analytical results for Pr in OREAS 250c (Certified Value 8.51 ppm) | [19 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 8.03 | 8.951 | 8.5614 | 8.2 | 7.97 | 4.6464 | 8.46 | [11 empty] | 1
| 1 | 2 | 8.48 | 8.9026 | 8.8559 | 8.8 | 7.99 | 4.6092 | 8.22 | [11 empty] | 8
| 1 | 3 | 8.45 | 8.9469 | 9.3919 | 8.8 | 7.87 | 4.4708 | 8.53 | [11 empty] | 16
| 1 | 4 | 8.52 | 9.2705 | 8.8564 | 8.3 | 7.81 | 4.5805 | 8.44 | [11 empty] | 8.5113
| 1 | 5 | 8.22 | 9.1571 | 9.6408 | 8.1 | 8.06 | 4.2945 | 8.08 | [11 empty] | 50
| 1 | 6 | 7.98 | 9.1598 | 8.5722 | 8.1 | 8.18 | 4.5405 | 8.55
| Mean |  | 8.28 | 9.0647 | 8.9798 | 8.3833 | 7.98 | 4.5237 | 8.38
| Median |  | 8.335 | 9.0541 | 8.8561 | 8.25 | 7.98 | 4.5605 | 8.45
| Std Dev. |  | 0.23774 | 0.15033 | 0.44239 | 0.33116 | 0.13236 | 0.12743 | 0.18815
| Rel.Std.Dev. |  | 2.871% | 1.658% | 4.927% | 3.95% | 1.659% | 2.817% | 2.245%
| PDM3 |  | -2.717% | 6.502% | 5.504% | -1.503% | -6.242% | -46.85% | -1.543%
| Z-Score (Absolute) |  | 0.17 | 1.15 | 1.01 | 0.01 | 0.67 | 6.5 | 0
| Analytical results for Rb in OREAS 250c (Certified Value 65 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 64.8 | 70.7329 | 65.1 | 58.9 | 68.1 | 64.6 | 58.7 | 45.756 | 78.7 | 64.6858 | 61.9 | 65.4 | 69.1 | 65.9 | 65.1 |  |  |  | 1
| 1 | 2 | 65.34 | 69.9124 | 64.4 | 61.7 | 68.9 | 63.7 | 58.8 | 43.0818 | 63.3 | 62.3601 | 61.1 | 67.2 | 69.7 | 65.6 | 65.7 |  |  |  | 32
| 1 | 3 | 67.63 | 70.6444 | 65 | 63.4 | 67.4 | 68.4 | 58.5 | 42.3522 | 64.9 | 63.6591 | 60.9 | 67.2 | 69.6 | 65.1 | 65.9 |  |  |  | 16
| 1 | 4 | 65.63 | 70.8477 | 65.5 | 58.8 | 67.1 | 66.1 | 56 | 45.0491 | 60 | 58.0253 | 61.6 | 69.9 | 68.9 | 64.8 | 64.6 |  |  |  | 64.7019
| 1 | 5 | 64.49 | 69.7461 | 65.2 | 58.7 | 68.1 | 65.9 | 59.4 | 37.5388 | 65.1 | 67.2973 | 59.6 | 68.9 | 68 | 66.9 | 61.8 |  |  |  | 51
| 1 | 6 | 61.52 | 70.5741 | 64.2 | 55 | 67.8 | 64.7 | 53.5 | 45.7775 | 64.7 | 57.384 | 61.6 | 70.2 | 69.4 | 66.7 | 63.4
| Mean |  | 64.9017 | 70.4096 | 64.9 | 59.4167 | 67.9 | 65.5667 | 57.4833 | 43.2592 | 66.1167 | 62.2353 | 61.1167 | 68.1333 | 69.1167 | 65.8333 | 64.4167
| Median |  | 65.07 | 70.6092 | 65.05 | 58.85 | 67.95 | 65.3 | 58.6 | 44.0654 | 64.8 | 63.0096 | 61.35 | 68.05 | 69.25 | 65.75 | 64.85
| Std Dev. |  | 1.9891 | 0.46176 | 0.498 | 2.8896 | 0.62929 | 1.6488 | 2.2798 | 3.1419 | 6.453 | 3.8712 | 0.82805 | 1.8544 | 0.62423 | 0.84301 | 1.5639
| Rel.Std.Dev. |  | 3.065% | 0.6558% | 0.7673% | 4.863% | 0.9268% | 2.515% | 3.966% | 7.263% | 9.76% | 6.22% | 1.355% | 2.722% | 0.9032% | 1.281% | 2.428%
| PDM3 |  | 0.3088% | 8.822% | 0.3062% | -8.169% | 4.943% | 1.337% | -11.16% | -33.14% | 2.187% | -3.812% | -5.541% | 5.303% | 6.823% | 1.749% | -0.4408%
| Z-Score (Absolute) |  | 0 | 1.24 | 0 | 1.23 | 0.67 | 0.15 | 1.67 | 4.87 | 0.27 | 0.6 | 0.85 | 0.73 | 0.95 | 0.21 | 0.11
| Analytical results for Re in OREAS 250c (Certified Value < 0.002 ppm) | [19 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | [5 empty] | 1
| 1 | 2 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | [5 empty] | 33
| 1 | 3 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | [5 empty] | 16
| 1 | 4 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | [5 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | [5 empty] | 52
| 1 | 6 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | < 0.003 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 250c (Certified Value 0.027 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0253 | 0.032827 | 0.43407 | 0.025 | 0.02 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.025783 | 0.03 | 0.02496 | 0.03 | 0.02 | 0.03 | 0.03 | 0.026 |  | 1
| 1 | 2 | 0.0253 | 0.031363 | 0.42415 | 0.024 | 0.03 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.023772 | 0.03 | 0.026578 | 0.03 | 0.02 | 0.04 | 0.03 | 0.026 |  | 18
| 1 | 3 | 0.0266 | 0.034364 | 0.43494 | 0.024 | 0.02 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.024069 | 0.03 | 0.026394 | 0.03 | 0.02 | 0.03 | 0.02 | 0.026 |  | 16
| 1 | 4 | 0.0246 | 0.031838 | 0.43393 | 0.025 | 0.02 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.025367 | 0.03 | 0.027832 | 0.03 | 0.02 | 0.03 | 0.03 | 0.027 |  | 0.027184
| 1 | 5 | 0.0258 | 0.029232 | 0.43449 | 0.025 | 0.03 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.023545 | 0.03 | 0.024965 | 0.03 | 0.02 | 0.03 | 0.03 | 0.026 |  | 53
| 1 | 6 | 0.0234 | 0.031955 | 0.43607 | 0.024 | 0.02 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.025211 | 0.03 | 0.027405 | 0.03 | 0.02 | 0.03 | 0.03 | 0.027
| Mean |  | 0.025167 | 0.03193 | 0.43294 | 0.0245 | 0.023333 | 0.03 | 0.03 | 0.04 |  | 0.024624 | 0.03 | 0.026356 | 0.03 | 0.02 | 0.031667 | 0.028333 | 0.026333
| Median |  | 0.0253 | 0.031896 | 0.43428 | 0.0245 | 0.02 | 0.03 | 0.03 | 0.04 |  | 0.02464 | 0.03 | 0.026486 | 0.03 | 0.02 | 0.03 | 0.03 | 0.026
| Std Dev. |  | 0.0010893 | 0.0016934 | 0.0043774 | 0.00054772 | 0.005164 | 0 | 0 | 0 |  | 0.00094212 | 0 | 0.0012012 | 0 | 0 | 0.0040825 | 0.0040825 | 0.0005164
| Rel.Std.Dev. |  | 4.329% | 5.304% | 1.011% | 2.236% | 22.13% | 0% | 0% | 0% |  | 3.826% | 0% | 4.558% | 0% | 0% | 12.89% | 14.41% | 1.961%
| PDM3 |  | -7.421% | 17.46% | 1493% | -9.874% | -14.17% | 10.36% | 10.36% | 47.15% |  | -9.416% | 10.36% | -3.048% | 10.36% | -26.43% | 16.49% | 4.228% | -3.129%
| Z-Score (Absolute) |  | 0.89 | 0.35 | 74.1 | 1.01 | 1.23 | 0 | 0 | 1.84 | 3.68 | 0.99 | 0 | 0.67 | 0 | 1.84 | 0.31 | 0.31 | 0.67
| Analytical results for Sb in OREAS 250c (Certified Value 0.88 ppm) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 0.93 | 0.84932 | < 0.1 | 0.9 | < 0.1 | 0.82 | < 5 | 0.92 | 0.9 | 0.9287 | 0.93 | 0.86695 | 0.85 | 0.81 | 0.93 | 0.86 | 1.02 |  | 1
| 1 | 2 | 0.9 | 0.80781 | 0.7734 | 0.9 | < 0.1 | 0.81 | < 5 | 0.89 | 0.9 | 1.0136 | 0.84 | 0.80087 | 0.82 | 0.85 | 0.98 | 0.82 | 1.03 |  | 34
| 1 | 3 | 0.92 | 0.85636 | 2.896 | 0.9 | < 0.1 | 0.83 | < 5 | 1 | 0.8 | 0.9803 | 0.85 | 0.80369 | 0.82 | 0.83 | 0.93 | 0.84 | 0.98 |  | 16
| 1 | 4 | 0.89 | 0.84677 | 1.622 | 0.9 | < 0.1 | 0.84 | < 5 | 0.95 | 0.9 | 0.9783 | 0.79 | 0.7494 | 0.85 | 0.88 | 0.92 | 0.82 | 1.01 |  | 0.88377
| 1 | 5 | 0.95 | 0.91544 | < 0.1 | 0.9 | < 0.1 | 0.82 | < 5 | 0.95 | 0.9 | 0.9456 | 0.83 | 0.8774 | 0.84 | 0.87 | 0.91 | 0.87 | 0.98 |  | 54
| 1 | 6 | 0.84 | 0.85238 | < 0.1 | 0.9 | < 0.1 | 0.81 | < 5 | 0.91 | 0.8 | 0.9715 | 0.88 | 0.7462 | 0.86 | 0.88 | 0.93 | 0.86 | 1.02
| Mean |  | 0.905 | 0.85468 | 1.7638 | 0.9 |  | 0.82167 |  | 0.93667 | 0.86667 | 0.96967 | 0.85333 | 0.80742 | 0.84 | 0.85333 | 0.93333 | 0.845 | 1.0067
| Median |  | 0.91 | 0.85085 | 1.622 | 0.9 |  | 0.82 |  | 0.935 | 0.9 | 0.9749 | 0.845 | 0.80228 | 0.845 | 0.86 | 0.93 | 0.85 | 1.015
| Std Dev. |  | 0.038341 | 0.034604 | 1.0684 | 0 |  | 0.01169 |  | 0.038816 | 0.05164 | 0.029592 | 0.04761 | 0.055877 | 0.016733 | 0.028752 | 0.024221 | 0.021679 | 0.021602
| Rel.Std.Dev. |  | 4.237% | 4.049% | 60.57% | 0% |  | 1.423% |  | 4.144% | 5.958% | 3.052% | 5.579% | 6.92% | 1.992% | 3.369% | 2.595% | 2.566% | 2.146%
| PDM3 |  | 2.402% | -3.292% | 99.58% | 1.836% |  | -7.027% |  | 5.985% | -1.936% | 9.719% | -3.444% | -8.64% | -4.953% | -3.444% | 5.608% | -4.387% | 13.91%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.7 | NA | 10.78 | 0.44 | 22.05 | 1.1 | NA | 1.54 | 0.02 | 0.63 | 0.2 | 0.02 | 1.05 | 0.13 | 2.04
Results from laboratories 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 250c (Certified Value 9.89 ppm) | [19 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 9.6 | 10.2013 | 9 | 11 | 10 | 9.6015 | 9.8 | 9 | 9.5121 | 8.7 | 10.276 | 9.8 | 10.2 | 10.2 | 9.7 | 11.1 |  |  | 1
| 1 | 2 | 9.9 | 10.2874 | 9 | 11 | 10.3 | 9.176 | 9.4 | 10 | 9.1839 | 8.4 | 10.0067 | 9.8 | 10.5 | 10.3 | 9.7 | 10.9 |  |  | 35
| 1 | 3 | 9.9 | 10.2535 | 9 | 11 | 10.2 | 8.7209 | 10.2 | 10 | 9.3536 | 8.4 | 9.8066 | 9.8 | 10.3 | 10.5 | 9.6 | 11.3 |  |  | 16
| 1 | 4 | 9.9 | 10.151 | 9 | 11 | 9.9 | 9.657 | 9.8 | 10 | 9.5718 | 7.8 | 10.1224 | 9.7 | 10.9 | 10.6 | 9.5 | 11.1 |  |  | 9.8918
| 1 | 5 | 9.6 | 10.2743 | 9 | 11 | 10.2 | 9.2944 | 9.8 | 10 | 9.1352 | 8.2 | 9.8211 | 9.4 | 10.6 | 10 | 10 | 10.7 |  |  | 55
| 1 | 6 | 8.9 | 10.3157 | 9 | 11 | 10.4 | 9.7236 | 9.8 | 10 | 9.5422 | 8.6 | 10.2714 | 9.7 | 10.8 | 10.6 | 9.9 | 10.5
| Mean |  | 9.6333 | 10.2472 | 9 | 11 | 10.1667 | 9.3622 | 9.8 | 9.8333 | 9.3831 | 8.35 | 10.0507 | 9.7 | 10.55 | 10.3667 | 9.7333 | 10.9333
| Median |  | 9.75 | 10.2639 | 9 | 11 | 10.2 | 9.448 | 9.8 | 10 | 9.4329 | 8.4 | 10.0646 | 9.75 | 10.55 | 10.4 | 9.7 | 11
| Std Dev. |  | 0.38816 | 0.060734 | 0 | 0 | 0.18619 | 0.38087 | 0.25298 | 0.40825 | 0.18951 | 0.32094 | 0.20921 | 0.15492 | 0.27386 | 0.24221 | 0.18619 | 0.29439
| Rel.Std.Dev. |  | 4.029% | 0.5927% | 0% | 0% | 1.831% | 4.068% | 2.581% | 4.152% | 2.02% | 3.844% | 2.082% | 1.597% | 2.596% | 2.336% | 1.913% | 2.693%
| PDM3 |  | -2.613% | 3.593% | -9.015% | 11.2% | 2.779% | -5.353% | -0.9279% | -0.5909% | -5.142% | -15.59% | 1.606% | -1.939% | 6.654% | 4.801% | -1.602% | 10.53%
| Z-Score (Absolute) |  | 0.27 | 0.72 | NA | NA | 0.59 | 0.71 | 0 | NA | 0.67 | 2.35 | 0.41 | 0.16 | 1.21 | 0.92 | 0.11 | 1.83
Results from laboratories 1.07, 1.09 and 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 250c (Indicative Value 0.76 ppm) | [19 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.8 | < 1.5 | 12.9833 | < 1 | 0.7 | < 1 | < 5 | 1 | < 1 | 0.0726 | < 1 | < 2 | 1.3 | < 1 | < 1 | 1 |  |  | 1
| 1 | 2 | 0.6 | < 1.5 | 16.1646 | < 1 | 0.6 | < 1 | < 5 | 1 | 1 | 0.091 | < 1 | < 2 | 1.3 | < 1 | < 1 | < 1 |  |  | 1
| 1 | 3 | 0.6 | < 1.5 | 15.6186 | < 1 | 0.6 | < 1 | < 5 | 1 | < 1 | 0.0842 | < 1 | < 2 | 1.4 | < 1 | < 1 | < 1 |  |  | 16
| 1 | 4 | 0.7 | < 1.5 | 14.333 | < 1 | 0.6 | < 1 | < 5 | 2 | < 1 | 0.0954 | < 1 | < 2 | 1.3 | < 1 | < 1 | < 1 |  |  | 0.75717
| 1 | 5 | < 0.5 | < 1.5 | 13.5937 | < 1 | 0.6 | < 1 | < 5 | 2 | < 1 | 0.0906 | < 1 | < 2 | 1.2 | < 1 | < 1 | 1 |  |  | 7
| 1 | 6 | 0.6 | < 1.5 | 10.2774 | < 1 | 0.4 | < 1 | < 5 | 1 | < 1 | 0.0899 | < 1 | < 2 | 1.3 | < 1 | < 1 | 1
| Mean |  | 0.66 |  | 13.8284 |  | 0.58333 |  |  | 1.3333 | 1 | 0.087283 |  |  | 1.3 |  |  | 1
| Median |  | 0.6 |  | 13.9634 |  | 0.6 |  |  | 1 | 1 | 0.09025 |  |  | 1.3 |  |  | 1
| Std Dev. |  | 0.089443 |  | 2.1117 |  | 0.098319 |  |  | 0.5164 |  | 0.008033 |  |  | 0.063246 |  |  | 0
| Rel.Std.Dev. |  | 13.55% |  | 15.27% |  | 16.85% |  |  | 38.73% |  | 9.203% |  |  | 4.865% |  |  | 0%
| PDM3 |  | -12.83% |  | 1726% |  | -22.96% |  |  | 76.09% | 32.07% | -88.47% |  |  | 71.69% |  |  | 32.07%
| Z-Score (Absolute) |  | 0.02 | 0.88 | 71.43 | 0.47 | 0.02 | 0.47 | 10.32 | 4.02 | 0.02 | 2.7 | 0.47 | 2.23 | 3.84 | 0.47 | 0.47 | 0.88
| Analytical results for Sm in OREAS 250c (Certified Value 6.27 ppm) | [19 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 5.91 | 6.3536 | 6.4438 | 5.7 | 5.83 | 6.5277 | 6.13 | [11 empty] | 1
| 1 | 2 | 6.2 | 6.4341 | 6.7304 | 6.3 | 5.95 | 6.2663 | 6.2 | [11 empty] | 5
| 1 | 3 | 6.18 | 6.4151 | 6.8332 | 7.1 | 5.86 | 6.3639 | 6.31 | [11 empty] | 16
| 1 | 4 | 6.33 | 6.7414 | 6.6434 | 6.2 | 5.62 | 6.5773 | 6.34 | [11 empty] | 6.2719
| 1 | 5 | 6.12 | 6.3918 | 6.8966 | 5.7 | 5.94 | 6.0651 | 5.99 | [11 empty] | 56
| 1 | 6 | 5.83 | 6.4046 | 6.5592 | 6.5 | 6.23 | 6.6144 | 6.03
| Mean |  | 6.095 | 6.4568 | 6.6844 | 6.25 | 5.905 | 6.4024 | 6.1667
| Median |  | 6.15 | 6.4098 | 6.6869 | 6.25 | 5.9 | 6.4458 | 6.165
| Std Dev. |  | 0.18897 | 0.14202 | 0.17002 | 0.5282 | 0.19887 | 0.21214 | 0.14348
| Rel.Std.Dev. |  | 3.1% | 2.2% | 2.543% | 8.451% | 3.368% | 3.314% | 2.327%
| PDM3 |  | -2.821% | 2.947% | 6.577% | -0.3493% | -5.85% | 2.081% | -1.678%
| Z-Score (Absolute) |  | 0.67 | 0.9 | 1.89 | 0 | 1.5 | 0.66 | 0.36
| Analytical results for Sn in OREAS 250c (Certified Value 3.59 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 3.6 | 3.9566 | 38.5257 | 3.9 | 1 | 3.3 | < 10 | 3.6 | 3.4 | 3.4746 | 3.7 | 3.5476 | 3.6 | 3.4 | 3.5 | 3.4 | 3.7 |  | 1
| 1 | 2 | 3.8 | 3.6704 | 44.1793 | 3.8 | < 1 | 3.3 | < 10 | 3.6 | 3.4 | 3.539 | 3.5 | 3.4331 | 3.6 | 3.5 | 3.6 | 3.4 | 3.8 |  | 21
| 1 | 3 | 3.6 | 3.6754 | 40.6233 | 3.9 | < 1 | 3.4 | < 10 | 3.7 | 3.6 | 3.6409 | 3.7 | 3.3665 | 3.6 | 3.4 | 3.6 | 3.5 | 3.9 |  | 16
| 1 | 4 | 3.8 | 3.6611 | 41.141 | 3.8 | < 1 | 3.2 | < 10 | 3.5 | 3.6 | 3.6433 | 3.4 | 3.386 | 3.6 | 3.5 | 3.6 | 3.4 | 3.9 |  | 3.5909
| 1 | 5 | 3.5 | 3.893 | 38.6557 | 3.9 | 1 | 3.4 | < 10 | 3.6 | 3.6 | 3.4559 | 3.5 | 3.3709 | 3.6 | 3.5 | 3.5 | 3.7 | 3.9 |  | 57
| 1 | 6 | 3.7 | 3.4041 | 35.1803 | 3.9 | 1 | 3.6 | < 10 | 4.1 | 3.5 | 3.5852 | 3.6 | 3.5309 | 3.6 | 3.6 | 3.6 | 3.7 | 3.7
| Mean |  | 3.6667 | 3.7101 | 39.7175 | 3.8667 | 1 | 3.3667 |  | 3.6833 | 3.5167 | 3.5565 | 3.5667 | 3.4392 | 3.6 | 3.4833 | 3.5667 | 3.5167 | 3.8167
| Median |  | 3.65 | 3.6729 | 39.6395 | 3.9 | 1 | 3.35 |  | 3.6 | 3.55 | 3.5621 | 3.55 | 3.4096 | 3.6 | 3.5 | 3.6 | 3.45 | 3.85
| Std Dev. |  | 0.12111 | 0.19649 | 3.0295 | 0.05164 | 0 | 0.13663 |  | 0.2137 | 0.098319 | 0.080766 | 0.12111 | 0.081201 | 0 | 0.075277 | 0.05164 | 0.1472 | 0.098319
| Rel.Std.Dev. |  | 3.303% | 5.296% | 7.628% | 1.336% | 0% | 4.058% |  | 5.802% | 2.796% | 2.271% | 3.395% | 2.361% | 0% | 2.161% | 1.448% | 4.186% | 2.576%
| PDM3 |  | 2.11% | 3.32% | 1006% | 7.68% | -72.15% | -6.244% |  | 2.574% | -2.067% | -0.9582% | -0.6746% | -4.225% | 0.2537% | -2.995% | -0.6746% | -2.067% | 6.287%
| Z-Score (Absolute) |  | 0.56 | 0.85 | 243.66 | 1.91 | NA | 1.46 | 9.55 | 0.67 | 0.45 | 0.18 | 0.11 | 0.97 | 0.11 | 0.67 | 0.11 | 0.45 | 1.57
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 250c (Certified Value 311 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 306.67 | 324.0972 | 315.7599 | 425.0299 | 283 | 289 | 307 | 327 | 302 | 311 | 307.8553 | 318 | 311.6868 | 318 | 303 | 315 | 309 | 281 | 1
| 1 | 2 | 312.21 | 338.2884 | 314.2662 | 424.1644 | 288 | 321 | 305 | 316 | 298 | 316 | 308.2956 | 309 | 299.8947 | 311 | 304 | 321 | 311 | 282 | 22
| 1 | 3 | 318.41 | 328.8492 | 312.5039 | 418.0125 | 289 | 322 | 309 | 321 | 315 | 309 | 314.2739 | 316 | 307.2161 | 311 | 306 | 309 | 304 | 279 | 16
| 1 | 4 | 316.83 | 335.664 | 313.4804 | 416.4947 | 291 | 302 | 306 | 325 | 308 | 313 | 320.8629 | 309 | 301.6424 | 315 | 314 | 320 | 303 | 286 | 311.4662
| 1 | 5 | 310.61 | 333.096 | 312.7444 | 423.1297 | 278 | 290 | 311 | 323 | 304 | 317 | 296.9176 | 319 | 305.7518 | 303 | 314 | 316 | 324 | 279 | 58
| 1 | 6 | 294.81 | 343.6452 | 314.4242 | 423.9788 | 280 | 281 | 310 | 326 | 308 | 315 | 314.2008 | 317 | 314.0448 | 313 | 318 | 325 | 312 | 278
| Mean |  | 309.9233 | 333.94 | 313.8632 | 421.8017 | 284.8333 | 300.8333 | 308 | 323 | 305.8333 | 313.5 | 310.401 | 314.6667 | 306.7061 | 311.8333 | 309.8333 | 317.6667 | 310.5 | 280.8333
| Median |  | 311.41 | 334.38 | 313.8733 | 423.5543 | 285.5 | 296 | 308 | 324 | 306 | 314 | 311.2482 | 316.5 | 306.4839 | 312 | 310 | 318 | 310 | 280
| Std Dev. |  | 8.5372 | 6.9222 | 1.21 | 3.6064 | 5.2694 | 17.3599 | 2.3664 | 4.0497 | 5.8793 | 3.0822 | 8.1438 | 4.5019 | 5.5113 | 5.0761 | 6.2743 | 5.5737 | 7.5565 | 2.9269
| Rel.Std.Dev. |  | 2.755% | 2.073% | 0.3855% | 0.855% | 1.85% | 5.771% | 0.7683% | 1.254% | 1.922% | 0.9832% | 2.624% | 1.431% | 1.797% | 1.628% | 2.025% | 1.755% | 2.434% | 1.042%
| PDM3 |  | -0.4954% | 7.215% | 0.7696% | 35.42% | -8.551% | -3.414% | -1.113% | 3.703% | -1.809% | 0.653% | -0.342% | 1.028% | -1.528% | 0.1179% | -0.5243% | 1.991% | -0.3102% | -9.835%
| Z-Score (Absolute) |  | 0.09 | 3.98 | 0.58 | 18.86 | 4.34 | 1.63 | 0.42 | 2.13 | 0.78 | 0.52 | 0.01 | 0.71 | 0.63 | 0.23 | 0.1 | 1.22 | 0.01 | 5.02
| Analytical results for Ta in OREAS 250c (Certified Value 2.07 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.09 | 2.2107 | 1.89 | < 0.1 | < 0.05 | 6.91 | 2.11 | 1.9 | 2.8657 | 2 | 2.0376 | 1.9 | 2 | 2.23 | 2.08 | 2.11 |  |  | 1
| 1 | 2 | 2.28 | 2.3622 | 1.94 | < 0.1 | < 0.05 | 7.1 | 2.05 | 2 | 3.0933 | 1.95 | 2.0439 | 1.8 | 2.04 | 2.33 | 2.03 | 2.05 |  |  | 23
| 1 | 3 | 2.55 | 2.8259 | 1.89 | < 0.1 | < 0.05 | 6.5667 | 2.19 | 1.9 | 2.9353 | 2.05 | 1.9671 | 1.8 | 2.04 | 2.37 | 2.09 | 2.11 |  |  | 16
| 1 | 4 | 2.28 | 2.7814 | 1.97 | < 0.1 | < 0.05 | 7.1367 | 2.12 | 1.9 | 2.9665 | 1.82 | 1.792 | 1.9 | 2.05 | 2.3 | 2.02 | 2.05 |  |  | 2.0748
| 1 | 5 | 2.16 | 2.6658 | 1.84 | < 0.1 | < 0.05 | 6.7733 | 2.1 | 1.9 | 2.8228 | 1.91 | 2.1472 | 1.8 | 2.16 | 2.26 | 2.14 | 2.11 |  |  | 59
| 1 | 6 | 2.04 | 2.8358 | 1.92 | < 0.1 | < 0.05 | 7.5267 | 2.13 | 1.9 | 2.9046 | 2 | 1.8322 | 1.9 | 2.15 | 2.35 | 2.08 | 2.06
| Mean |  | 2.2333 | 2.6136 | 1.9083 |  |  | 7.0022 | 2.1167 | 1.9167 | 2.9314 | 1.955 | 1.97 | 1.85 | 2.0733 | 2.3067 | 2.0733 | 2.0817
| Median |  | 2.22 | 2.7236 | 1.905 |  |  | 7.005 | 2.115 | 1.9 | 2.92 | 1.975 | 2.0023 | 1.85 | 2.045 | 2.315 | 2.08 | 2.085
| Std Dev. |  | 0.18327 | 0.2649 | 0.04535 |  |  | 0.33265 | 0.045461 | 0.040825 | 0.094114 | 0.08167 | 0.13575 | 0.054772 | 0.065625 | 0.053914 | 0.043665 | 0.031252
| Rel.Std.Dev. |  | 8.206% | 10.14% | 2.376% |  |  | 4.751% | 2.148% | 2.13% | 3.211% | 4.177% | 6.891% | 2.961% | 3.165% | 2.337% | 2.106% | 1.501%
| PDM3 |  | 7.64% | 25.97% | -8.024% |  |  | 237.5% | 2.017% | -7.623% | 41.28% | -5.775% | -5.052% | -10.84% | -0.07198% | 11.17% | -0.07198% | 0.3297%
| Z-Score (Absolute) |  | 0.66 | 2.24 | 0.68 | 8.4 | 8.5 | 20.45 | 0.18 | 0.65 | 3.56 | 0.49 | 0.43 | 0.93 | 0 | 0.97 | 0 | 0.03
| Analytical results for Tb in OREAS 250c (Certified Value 0.69 ppm) | [19 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.17 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 0.68 | 0.81485 | 0.66457 | 0.7 | 0.71 | 0.69296 | 0.8 | 0.68 | [10 empty] | 1
| 1 | 2 | 0.7 | 0.76614 | 0.66713 | 0.7 | 0.69 | 0.6954 | 0.8 | 0.66 | [10 empty] | 24
| 1 | 3 | 0.7 | 0.76084 | 0.69588 | 0.7 | 0.68 | 0.69611 | 0.7 | 0.68 | [10 empty] | 16
| 1 | 4 | 0.71 | 0.80644 | 0.69462 | 0.7 | 0.7 | 0.69635 | 0.8 | 0.7 | [10 empty] | 0.68798
| 1 | 5 | 0.68 | 0.80315 | 0.69468 | 0.6 | 0.69 | 0.69935 | 0.7 | 0.65 | [10 empty] | 60
| 1 | 6 | 0.71 | 0.76213 | 0.65679 | 0.6 | 0.72 | 0.69568 | 0.7 | 0.65
| Mean |  | 0.69667 | 0.78559 | 0.67895 | 0.66667 | 0.69833 | 0.69597 | 0.75 | 0.67
| Median |  | 0.7 | 0.78465 | 0.68087 | 0.7 | 0.695 | 0.69589 | 0.75 | 0.67
| Std Dev. |  | 0.013663 | 0.02506 | 0.017985 | 0.05164 | 0.01472 | 0.002052 | 0.054772 | 0.02
| Rel.Std.Dev. |  | 1.961% | 3.19% | 2.649% | 7.746% | 2.108% | 0.2948% | 7.303% | 2.985%
| PDM3 |  | 1.262% | 14.19% | -1.314% | -3.098% | 1.504% | 1.161% | 9.014% | -2.614%
| Z-Score (Absolute) |  | 0.02 | 6.21 | 1.21 | NA | 0.14 | 0.02 | NA | 1.83
Results from laboratories 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 250c (Indicative Value 0.029 ppm) | [19 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | < 0.2 | 0.10127 | 0.1 | < 0.1 | < 0.05 | 7.208 | < 0.05 | < 0.5 | 0.0434 | 0.05 | 0.17122 | 0.03 | < 0.05 | < 0.05 | < 0.05 |  |  |  | 1
| 1 | 2 | < 0.2 | 0.11026 | 0.1 | < 0.1 | < 0.05 | 7.4293 | 0.09 | < 0.5 | < 0.03 | 0.05 | 0.17237 | 0.05 | 0.06 | < 0.05 | < 0.05 |  |  |  | 2
| 1 | 3 | < 0.2 | < 0.1 | 0.1 | < 0.1 | < 0.05 | 7.1253 | < 0.05 | < 0.5 | < 0.03 | < 0.05 | 0.13008 | 0.04 | 0.05 | < 0.05 | < 0.05 |  |  |  | 16
| 1 | 4 | < 0.2 | 0.13038 | 0.1 | < 0.1 | < 0.05 | 8.448 | < 0.05 | < 0.5 | < 0.03 | < 0.05 | 0.14203 | 0.04 | < 0.05 | < 0.05 | < 0.05 |  |  |  | 0.028842
| 1 | 5 | < 0.2 | < 0.1 | 0.1 | < 0.1 | < 0.05 | 7.008 | < 0.05 | < 0.5 | 0.0386 | < 0.05 | 0.15935 | 0.03 | < 0.05 | < 0.05 | < 0.05 |  |  |  | 8
| 1 | 6 | < 0.2 | 0.12251 | 0.1 | < 0.1 | < 0.05 | 8.0533 | < 0.05 | < 0.5 | < 0.03 | < 0.05 | 0.14305 | 0.05 | < 0.05 | 0.05 | < 0.05
| Mean |  |  | 0.1161 | 0.1 |  |  | 7.5453 | 0.09 |  | 0.041 | 0.05 | 0.15302 | 0.04 | 0.055 | 0.05
| Median |  |  | 0.11638 | 0.1 |  |  | 7.3187 | 0.09 |  | 0.041 | 0.05 | 0.1512 | 0.04 | 0.055 | 0.05
| Std Dev. |  |  | 0.012895 | 1.5202e-17 |  |  | 0.5771 |  |  | 0.0033941 | 0 | 0.017275 | 0.0089443 | 0.0070711
| Rel.Std.Dev. |  |  | 11.11% | 1.52e-14% |  |  | 7.648% |  |  | 8.278% | 0% | 11.29% | 22.36% | 12.86%
| PDM3 |  |  | 302.5% | 246.7% |  |  | 26061% | 212% |  | 42.15% | 73.36% | 430.5% | 38.69% | 90.69% | 73.36%
| Z-Score (Absolute) |  | 2.7 | 2.43 | 2.7 | 0.45 | 0.67 | 337.39 | 0.19 | 9.44 | 0.73 | 0.3 | 5.08 | 0 | 0.22 | 0.49 | 0.67
| Analytical results for Th in OREAS 250c (Certified Value 8.68 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 8.34 | 8.8986 | 8.6909 | 8.57 | 9.5 | 7.28 | 8.62 | 8.6 | 11.37 | 8.64 | 8.7765 | 8.77 | 8.6 | 9.15 | 8.41 | 8.75 |  |  | 1
| 1 | 2 | 8.81 | 8.9151 | 8.4818 | 8.55 | 10 | 7.34 | 8.56 | 8.5 | 11.5667 | 8.05 | 8.684 | 8.72 | 9.32 | 9.4 | 8.55 | 8.7 |  |  | 26
| 1 | 3 | 8.47 | 8.7515 | 8.6377 | 8.74 | 10.2 | 7.55 | 8.97 | 8.7 | 11.6984 | 8.33 | 8.5528 | 8.37 | 9.28 | 8.99 | 8.47 | 8.67 |  |  | 16
| 1 | 4 | 8.91 | 8.9884 | 8.8618 | 8.77 | 9.4 | 7.31 | 8.64 | 8.7 | 11.8806 | 7.83 | 7.6878 | 8.11 | 9.12 | 9.16 | 8.5 | 8.81 |  |  | 8.6755
| 1 | 5 | 8.45 | 8.7685 | 8.3572 | 8.56 | 9.4 | 7.45 | 8.84 | 8.7 | 11.3059 | 7.97 | 9.2648 | 8.81 | 9.33 | 8.84 | 8.92 | 8.44 |  |  | 61
| 1 | 6 | 8.13 | 8.9798 | 8.4588 | 8.77 | 9 | 7.32 | 8.69 | 8.3 | 11.4897 | 8.29 | 7.7964 | 8.76 | 8.89 | 9.26 | 9 | 8.46
| Mean |  | 8.5183 | 8.8836 | 8.5814 | 8.66 | 9.5833 | 7.375 | 8.72 | 8.5833 | 11.5519 | 8.185 | 8.4604 | 8.59 | 9.09 | 9.1333 | 8.6417 | 8.6383
| Median |  | 8.46 | 8.9068 | 8.5597 | 8.655 | 9.45 | 7.33 | 8.665 | 8.65 | 11.5282 | 8.17 | 8.6184 | 8.74 | 9.2 | 9.155 | 8.525 | 8.685
| Std Dev. |  | 0.2926 | 0.10213 | 0.18385 | 0.11027 | 0.44008 | 0.10368 | 0.15479 | 0.16021 | 0.21317 | 0.2931 | 0.60716 | 0.28475 | 0.29244 | 0.19735 | 0.25199 | 0.15355
| Rel.Std.Dev. |  | 3.435% | 1.15% | 2.142% | 1.273% | 4.592% | 1.406% | 1.775% | 1.867% | 1.845% | 3.581% | 7.177% | 3.315% | 3.217% | 2.161% | 2.916% | 1.778%
| PDM3 |  | -1.812% | 2.399% | -1.085% | -0.1786% | 10.46% | -14.99% | 0.513% | -1.062% | 33.16% | -5.654% | -2.479% | -0.9854% | 4.778% | 5.277% | -0.3899% | -0.4283%
| Z-Score (Absolute) |  | 0.54 | 1.09 | 0.26 | 0.09 | 4.22 | 5.66 | 0.36 | 0.25 | 13.03 | 2.04 | 0.8 | 0.22 | 2.01 | 2.21 | 0.01 | 0.01
| Analytical results for Ti in OREAS 250c (Certified Value 0.567 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.5797 | 0.59124 | 0.57127 | 0.49463 | 0.558 | 0.482 | 0.556 | 0.55 | 0.552 | 0.56 | 0.44694 | 0.579 | 0.59963 | 0.57 | 0.564 | 0.563 | 0.559 | 0.502 | 1
| 1 | 2 | 0.5691 | 0.59373 | 0.57484 | 0.48724 | 0.554 | 0.185 | 0.549 | 0.54 | 0.543 | 0.577 | 0.42602 | 0.56 | 0.62057 | 0.56 | 0.565 | 0.572 | 0.556 | 0.488 | 27
| 1 | 3 | 0.5704 | 0.59515 | 0.57806 | 0.49439 | 0.567 | 0.165 | 0.548 | 0.53 | 0.577 | 0.563 | 0.43258 | 0.572 | 0.60642 | 0.56 | 0.569 | 0.544 | 0.555 | 0.481 | 16
| 1 | 4 | 0.5762 | 0.59508 | 0.56966 | 0.49288 | 0.56 | 0.27 | 0.541 | 0.56 | 0.564 | 0.574 | 0.44842 | 0.563 | 0.62143 | 0.56 | 0.583 | 0.562 | 0.562 | 0.512 | 0.56692
| 1 | 5 | 0.5859 | 0.59662 | 0.56962 | < 0.1 | 0.562 | 0.414 | 0.559 | 0.54 | 0.56 | 0.578 | 0.41871 | 0.576 | 0.59994 | 0.56 | 0.578 | 0.564 | 0.579 | 0.495 | 62
| 1 | 6 | 0.5844 | 0.5937 | 0.57203 | < 0.1 | 0.562 | 0.462 | 0.568 | 0.57 | 0.558 | 0.57 | 0.44285 | 0.579 | 0.628 | 0.57 | 0.588 | 0.586 | 0.565 | 0.487
| Mean |  | 0.57762 | 0.59425 | 0.57258 | 0.49228 | 0.5605 | 0.32967 | 0.5535 | 0.54833 | 0.559 | 0.57033 | 0.43592 | 0.5715 | 0.61266 | 0.56333 | 0.5745 | 0.56517 | 0.56267 | 0.49417
| Median |  | 0.57795 | 0.59441 | 0.57165 | 0.49364 | 0.561 | 0.342 | 0.5525 | 0.545 | 0.559 | 0.572 | 0.43771 | 0.574 | 0.61349 | 0.56 | 0.5735 | 0.5635 | 0.5605 | 0.4915
| Std Dev. |  | 0.0070084 | 0.0018274 | 0.0032986 | 0.0034525 | 0.0043704 | 0.14102 | 0.0095237 | 0.01472 | 0.011454 | 0.0074476 | 0.012091 | 0.0082158 | 0.01221 | 0.005164 | 0.009975 | 0.013747 | 0.0088242 | 0.011339
| Rel.Std.Dev. |  | 1.213% | 0.3075% | 0.5761% | 0.7013% | 0.7797% | 42.78% | 1.721% | 2.684% | 2.049% | 1.306% | 2.774% | 1.438% | 1.993% | 0.9167% | 1.736% | 2.432% | 1.568% | 2.295%
| PDM3 |  | 1.886% | 4.82% | 0.9977% | -13.17% | -1.133% | -41.85% | -2.368% | -3.279% | -1.398% | 0.6013% | -23.11% | 0.8071% | 8.068% | -0.6334% | 1.336% | -0.31% | -0.751% | -12.83%
| Z-Score (Absolute) |  | 0.94 | 2 | 0.61 | 13.96 | 0.16 | 14.93 | 0.61 | 0.94 | 0.26 | 0.47 | 8.13 | 0.54 | 3.18 | 0.02 | 0.74 | 0.14 | 0.02 | 4.4
| Analytical results for Tl in OREAS 250c (Certified Value 0.35 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.36 | 0.11514 | 0.34976 | 0.35 | 0.37 | < 0.02 | 0.32 | < 0.5 | < 0.4821 | 0.33 | 0.37294 | 0.38 | 0.33 | 0.34 | 0.35 | 0.37 |  |  | 1
| 1 | 2 | 0.39 | 0.10288 | 0.33919 | 0.34 | 0.4 | < 0.02 | 0.32 | < 0.5 | < 0.4249 | 0.33 | 0.34298 | 0.37 | 0.34 | 0.35 | 0.35 | 0.38 |  |  | 28
| 1 | 3 | 0.35 | 0.11 | 0.3532 | 0.35 | 0.39 | < 0.02 | 0.34 | < 0.5 | < 0.5785 | 0.35 | 0.37065 | 0.36 | 0.34 | 0.35 | 0.3 | 0.37 |  |  | 16
| 1 | 4 | 0.37 | 0.10582 | 0.35375 | 0.34 | 0.38 | < 0.02 | 0.34 | < 0.5 | < 0.5277 | 0.32 | 0.28398 | 0.37 | 0.34 | 0.34 | 0.36 | 0.37 |  |  | 0.35326
| 1 | 5 | 0.36 | 0.10093 | 0.36776 | 0.34 | 0.36 | < 0.02 | 0.34 | < 0.5 | < 0.5935 | 0.34 | 0.35244 | 0.37 | 0.35 | 0.34 | 0.34 | 0.36 |  |  | 63
| 1 | 6 | 0.34 | < 0.1 | 0.34885 | 0.34 | 0.36 | < 0.02 | 0.33 | < 0.5 | < 0.4339 | 0.33 | 0.28233 | 0.37 | 0.34 | 0.35 | 0.37 | 0.38
| Mean |  | 0.36167 | 0.10695 | 0.35208 | 0.34333 | 0.37667 |  | 0.33167 |  |  | 0.33333 | 0.33422 | 0.37 | 0.34 | 0.345 | 0.345 | 0.37167
| Median |  | 0.36 | 0.10582 | 0.35148 | 0.34 | 0.375 |  | 0.335 |  |  | 0.33 | 0.34771 | 0.37 | 0.34 | 0.345 | 0.35 | 0.37
| Std Dev. |  | 0.017224 | 0.0057124 | 0.0092959 | 0.005164 | 0.01633 |  | 0.0098319 |  |  | 0.010328 | 0.041115 | 0.0063246 | 0.0063246 | 0.0054772 | 0.02429 | 0.0075277
| Rel.Std.Dev. |  | 4.762% | 5.341% | 2.64% | 1.504% | 4.335% |  | 2.964% |  |  | 3.098% | 12.3% | 1.709% | 1.86% | 1.588% | 7.041% | 2.025%
| PDM3 |  | 2.379% | -69.72% | -0.3343% | -2.811% | 6.625% |  | -6.114% |  |  | -5.642% | -5.391% | 4.738% | -3.755% | -2.339% | -2.339% | 5.209%
| Z-Score (Absolute) |  | 0.89 | 10.83 | 0.46 | 0.07 | 1.55 | 14.71 | 0.44 | 4.06 | 3.91 | 0.37 | 0.33 | 1.26 | 0.07 | 0.15 | 0.15 | 1.33
| Analytical results for Tm in OREAS 250c (Certified Value 0.2 ppm) | [19 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.22 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.2 | 0.22836 | 0.18789 | 0.2 | 0.19 | 0.21152 | 0.19 | [11 empty] | 1
| 1 | 2 | 0.24 | 0.21216 | 0.19295 | 0.2 | 0.19 | 0.2082 | 0.18 | [11 empty] | 7
| 1 | 3 | 0.21 | 0.21922 | 0.19278 | 0.2 | 0.19 | 0.21 | 0.19 | [11 empty] | 16
| 1 | 4 | 0.21 | 0.21939 | 0.18992 | 0.2 | 0.19 | 0.2142 | 0.2 | [11 empty] | 0.20127
| 1 | 5 | 0.2 | 0.21595 | 0.19141 | 0.2 | 0.2 | 0.21168 | 0.19 | [11 empty] | 64
| 1 | 6 | 0.2 | 0.21837 | 0.18691 | 0.2 | 0.21 | 0.211 | 0.18
| Mean |  | 0.21 | 0.21891 | 0.19031 | 0.2 | 0.195 | 0.2111 | 0.18833
| Median |  | 0.205 | 0.21879 | 0.19066 | 0.2 | 0.19 | 0.21126 | 0.19
| Std Dev. |  | 0.015492 | 0.005372 | 0.0025232 | 3.0405e-17 | 0.0083666 | 0.0019876 | 0.0075277
| Rel.Std.Dev. |  | 7.377% | 2.454% | 1.326% | 1.52e-14% | 4.291% | 0.9415% | 3.997%
| PDM3 |  | 4.335% | 8.76% | -5.448% | -0.6333% | -3.117% | 4.882% | -6.43%
| Z-Score (Absolute) |  | 0.49 | 1.06 | 0.79 | NA | 0.49 | 0.56 | 0.92
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 250c (Certified Value 1.62 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.63 | 1.6697 | 1.5151 | 1.69 | 1.8 | 1.6 | 1.6 | 1.5 | 1.7824 | 1.5 | 1.6344 | 1.73 | 1.5 | 1.6 | 1.5 | 1.83 |  |  | 1
| 1 | 2 | 1.72 | 1.6145 | 1.635 | 1.61 | 1.9 | 1.6 | 1.6 | 1.6 | 1.8126 | 1.4 | 1.6497 | 1.69 | 1.6 | 1.7 | 1.6 | 1.79 |  |  | 30
| 1 | 3 | 1.71 | 1.6176 | 1.4339 | 1.65 | 1.9 | 1.5 | 1.7 | 1.5 | 1.8354 | 1.5 | 1.5698 | 1.67 | 1.5 | 1.6 | 1.6 | 1.84 |  |  | 16
| 1 | 4 | 1.76 | 1.6977 | 1.5519 | 1.65 | 1.8 | 1.6 | 1.6 | 1.5 | 1.8324 | 1.4 | 1.4783 | 1.7 | 1.6 | 1.6 | 1.6 | 1.79 |  |  | 1.6213
| 1 | 5 | 1.66 | 1.639 | 1.4623 | 1.62 | 1.7 | 1.6 | 1.6 | 1.5 | 1.8391 | 1.5 | 1.706 | 1.67 | 1.7 | 1.6 | 1.6 | 1.78 |  |  | 65
| 1 | 6 | 1.6 | 1.64 | 1.5024 | 1.63 | 1.7 | 1.6 | 1.6 | 1.5 | 1.8189 | 1.4 | 1.4621 | 1.65 | 1.6 | 1.7 | 1.6 | 1.77
| Mean |  | 1.68 | 1.6464 | 1.5168 | 1.6417 | 1.8 | 1.5833 | 1.6167 | 1.5167 | 1.8201 | 1.45 | 1.5834 | 1.685 | 1.5833 | 1.6333 | 1.5833 | 1.8
| Median |  | 1.685 | 1.6395 | 1.5087 | 1.64 | 1.8 | 1.6 | 1.6 | 1.5 | 1.8256 | 1.45 | 1.6021 | 1.68 | 1.6 | 1.6 | 1.6 | 1.79
| Std Dev. |  | 0.060332 | 0.031991 | 0.071072 | 0.028577 | 0.089443 | 0.040825 | 0.040825 | 0.040825 | 0.021097 | 0.054772 | 0.097944 | 0.028107 | 0.075277 | 0.05164 | 0.040825 | 0.028284
| Rel.Std.Dev. |  | 3.591% | 1.943% | 4.686% | 1.741% | 4.969% | 2.578% | 2.525% | 2.692% | 1.159% | 3.777% | 6.186% | 1.668% | 4.754% | 3.162% | 2.578% | 1.571%
| PDM3 |  | 3.619% | 1.548% | -6.449% | 1.254% | 11.02% | -2.343% | -0.2874% | -6.455% | 12.26% | -10.57% | -2.34% | 3.927% | -2.343% | 0.7405% | -2.343% | 11.02%
| Z-Score (Absolute) |  | 0.77 | 0.3 | 1.51 | 0.23 | 2.44 | 0.58 | 0.12 | 1.51 | 2.72 | 2.44 | 0.58 | 0.84 | 0.58 | 0.12 | 0.58 | 2.44
| Analytical results for V in OREAS 250c (Certified Value 70 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 70 | 69.5432 | 71.4922 | 58.0928 | 66 | 53 | 69 | 72.2667 | 68 | 73 | 61.5772 | 71 | 72.3319 | 69 | 69 | 69 | 70 | 69 | 1
| 1 | 2 | 70 | 67.326 | 71.7887 | 58.2228 | 67 | 33 | 68 | 71.025 | 69 | 75 | 62.057 | 69 | 73.6986 | 68 | 70 | 70 | 69 | 66 | 31
| 1 | 3 | 72 | 66.8458 | 71.7824 | 56.6389 | 65 | 32 | 69 | 71.54 | 71 | 74 | 61.5552 | 70 | 70.6072 | 68 | 70 | 69 | 69 | 66 | 16
| 1 | 4 | 70 | 66.4048 | 71.4833 | 56.4924 | 67 | 35 | 68 | 73.1367 | 69 | 75 | 61.9162 | 69 | 73.4608 | 68 | 71 | 70 | 70 | 67 | 69.932
| 1 | 5 | 75 | 73.0296 | 71.6892 | 58.0647 | 65 | 47 | 70 | 70.3333 | 69 | 76 | 61.9756 | 72 | 71.0314 | 66 | 71 | 70 | 72 | 65 | 66
| 1 | 6 | 66 | 69.7564 | 72.0572 | 58.5578 | 66 | 50 | 69 | 73.3433 | 68 | 76 | 61.9774 | 71 | 72.9063 | 68 | 72 | 71 | 71 | 65
| Mean |  | 70.5 | 68.8176 | 71.7155 | 57.6782 | 66 | 41.6667 | 68.8333 | 71.9408 | 69 | 74.8333 | 61.8431 | 70.3333 | 72.3394 | 67.8333 | 70.5 | 69.8333 | 70.1667 | 66.3333
| Median |  | 70 | 68.4346 | 71.7358 | 58.0787 | 66 | 41 | 69 | 71.9033 | 69 | 75 | 61.9459 | 70.5 | 72.6191 | 68 | 70.5 | 70 | 70 | 66
| Std Dev. |  | 2.9496 | 2.4933 | 0.21498 | 0.88066 | 0.89443 | 9.3737 | 0.75277 | 1.1905 | 1.0954 | 1.169 | 0.21923 | 1.2111 | 1.2757 | 0.98319 | 1.0488 | 0.75277 | 1.169 | 1.5055
| Rel.Std.Dev. |  | 4.184% | 3.623% | 0.2998% | 1.527% | 1.355% | 22.5% | 1.094% | 1.655% | 1.588% | 1.562% | 0.3545% | 1.722% | 1.764% | 1.449% | 1.488% | 1.078% | 1.666% | 2.27%
| PDM3 |  | 0.8122% | -1.594% | 2.55% | -17.52% | -5.623% | -40.42% | -1.571% | 2.873% | -1.333% | 7.009% | -11.57% | 0.5739% | 3.442% | -3.001% | 0.8122% | -0.1411% | 0.3356% | -5.146%
| Z-Score (Absolute) |  | 0.38 | 0.21 | 0.8 | 4.08 | 1.19 | 9.64 | 0.2 | 0.88 | 0.14 | 1.88 | 2.63 | 0.32 | 1.02 | 0.55 | 0.38 | 0.14 | 0.26 | 1.07
| Analytical results for W in OREAS 250c (Certified Value 2.05 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 2.1 | 2.1553 | 2.0283 | 6.3924 | 1.9 | 0.1 | 1.2 | < 10 | 2.1 | 1.8 | 2.189 | 2.1 | 1.9891 | 2.3 | 1.9 | 2.2 | 2.1 | 2 | 1
| 1 | 2 | 2.2 | 2.3671 | 2.055 | 4.1707 | 1.7 | < 0.1 | 1.2 | < 10 | 2 | 2 | 2.1456 | 2 | 1.9704 | 2.2 | 2 | 2.3 | 2 | 2 | 32
| 1 | 3 | 2.1 | 2.2138 | 1.9872 | 4.6433 | 1.8 | < 0.1 | 1.3 | < 10 | 2.1 | 1.9 | 2.251 | 2.1 | 1.864 | 2.3 | 1.9 | 2.2 | 2.1 | 1.9 | 16
| 1 | 4 | 2.2 | 2.1834 | 2.053 | 3.8136 | 1.8 | < 0.1 | 1.2 | < 10 | 2.1 | 1.8 | 2.2662 | 2 | 1.9994 | 2.1 | 2 | 2.2 | 2.1 | 1.9 | 2.0536
| 1 | 5 | 2.2 | 2.1383 | 1.988 | 4.3177 | 1.9 | < 0.1 | 1.2 | < 10 | 2 | 1.9 | 2.1147 | 2 | 2.0586 | 2.1 | 2.1 | 2.1 | 2.4 | 1.9 | 67
| 1 | 6 | 2 | 2.231 | 1.9367 | 5.1249 | 1.9 | < 0.1 | 1.3 | < 10 | 2.1 | 1.9 | 2.1673 | 2.1 | 1.8292 | 2.2 | 2.1 | 2.3 | 2.2 | 1.9
| Mean |  | 2.1333 | 2.2148 | 2.008 | 4.7438 | 1.8333 | 0.1 | 1.2333 |  | 2.0667 | 1.8833 | 2.189 | 2.05 | 1.9518 | 2.2 | 2 | 2.2167 | 2.15 | 1.9333
| Median |  | 2.15 | 2.1986 | 2.0082 | 4.4805 | 1.85 | 0.1 | 1.2 |  | 2.1 | 1.9 | 2.1782 | 2.05 | 1.9798 | 2.2 | 2 | 2.2 | 2.1 | 1.9
| Std Dev. |  | 0.08165 | 0.082282 | 0.04596 | 0.92179 | 0.08165 |  | 0.05164 |  | 0.05164 | 0.075277 | 0.059462 | 0.054772 | 0.087362 | 0.089443 | 0.089443 | 0.075277 | 0.13784 | 0.05164
| Rel.Std.Dev. |  | 3.827% | 3.715% | 2.289% | 19.43% | 4.454% |  | 4.187% |  | 2.499% | 3.997% | 2.716% | 2.672% | 4.476% | 4.066% | 4.472% | 3.396% | 6.411% | 2.671%
| PDM3 |  | 3.885% | 7.852% | -2.217% | 131% | -10.72% | -95.13% | -39.94% |  | 0.6384% | -8.289% | 6.594% | -0.1732% | -4.956% | 7.131% | -2.608% | 7.943% | 4.696% | -5.854%
| Z-Score (Absolute) |  | 0.37 | 0.77 | 0.25 | 13.3 | 1.11 | 9.91 | 4.09 | 14.57 | 0.04 | 0.87 | 0.65 | 0.04 | 0.53 | 0.7 | 0.29 | 0.78 | 0.45 | 0.62
| Analytical results for Y in OREAS 250c (Certified Value 15.7 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 16.27 | 17.0879 | 15.7418 | 15.8 | 15.5 | 15.2 | 14.8267 | 16.3 | 14.2 | 14.6685 | 16.6 | 16.0784 | 16.5 | 15.8 | 16.6 | 14.8 | 14.9 |  | 1
| 1 | 2 | 16.19 | 16.9114 | 15.8273 | 16.2 | 16.7 | 15.3 | 14.6 | 15.8 | 14.2 | 13.6889 | 15.9 | 15.5067 | 16.3 | 15.9 | 17 | 14.6 | 14.4 |  | 33
| 1 | 3 | 15.82 | 16.9913 | 15.6429 | 16.1 | 16.8 | 15.2 | 14.28 | 17.9 | 14.2 | 14.0188 | 16.6 | 15.2042 | 16.1 | 16.3 | 17.2 | 14.8 | 14.6 |  | 16
| 1 | 4 | 16.41 | 16.8959 | 15.9603 | 16 | 16.1 | 15.2 | 15.1833 | 16.2 | 14.1 | 14.3385 | 14.9 | 14.2922 | 16.2 | 16.6 | 16.8 | 15.4 | 14.9 |  | 15.6643
| 1 | 5 | 15.96 | 16.5183 | 16.2529 | 16.4 | 15.3 | 15.4 | 14.1933 | 16 | 14.7 | 13.4492 | 15.9 | 16.6133 | 15.7 | 20.3 | 16.3 | 16.6 | 14.7 |  | 68
| 1 | 6 | 15.18 | 16.8938 | 16.2677 | 15.7 | 15.2 | 15.9 | 15.2233 | 17.2 | 13.8 | 14.2985 | 16.2 | 15.0915 | 16.3 | 17 | 17.1 | 15.6 | 14
| Mean |  | 15.9717 | 16.8831 | 15.9488 | 16.0333 | 15.9333 | 15.3667 | 14.7178 | 16.5667 | 14.2 | 14.0771 | 16.0167 | 15.4644 | 16.1833 | 16.9833 | 16.8333 | 15.3 | 14.5833
| Median |  | 16.075 | 16.9037 | 15.8938 | 16.05 | 15.8 | 15.25 | 14.7133 | 16.25 | 14.2 | 14.1587 | 16.05 | 15.3555 | 16.25 | 16.45 | 16.9 | 15.1 | 14.65
| Std Dev. |  | 0.44242 | 0.19381 | 0.26287 | 0.2582 | 0.70616 | 0.27325 | 0.43926 | 0.81158 | 0.28983 | 0.45066 | 0.63061 | 0.8099 | 0.27142 | 1.6845 | 0.33862 | 0.74565 | 0.34303
| Rel.Std.Dev. |  | 2.77% | 1.148% | 1.648% | 1.61% | 4.432% | 1.778% | 2.985% | 4.899% | 2.041% | 3.201% | 3.937% | 5.237% | 1.677% | 9.919% | 2.012% | 4.874% | 2.352%
| PDM3 |  | 1.962% | 7.781% | 1.816% | 2.356% | 1.718% | -1.9% | -6.042% | 5.761% | -9.348% | -10.13% | 2.25% | -1.276% | 3.314% | 8.421% | 7.463% | -2.326% | -6.901%
| Z-Score (Absolute) |  | 0.02 | 1.02 | 0 | 0.09 | 0.02 | 0.64 | 1.34 | 0.67 | 1.91 | 2.04 | 0.07 | 0.53 | 0.26 | 1.13 | 0.97 | 0.71 | 1.49
| Analytical results for Yb in OREAS 250c (Certified Value 1.28 ppm) | [19 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.14 | 1.17 | 1.22 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 2
3
| 1 | 1 | 1.26 | 1.376 | 1.1728 | 1.4 | 1.26 | 1.3179 | 1.4 | 1.1 | [10 empty] | 1
| 1 | 2 | 1.32 | 1.3487 | 1.3362 | 1.5 | 1.24 | 1.2082 | 1.3 | 1.1 | [10 empty] | 34
| 1 | 3 | 1.32 | 1.313 | 1.4869 | 1.4 | 1.23 | 1.1982 | 1.2 | 1.1 | [10 empty] | 16
| 1 | 4 | 1.39 | 1.3861 | 1.2676 | 1.4 | 1.25 | 1.2779 | 1.3 | 1.2 | [10 empty] | 1.2817
| 1 | 5 | 1.31 | 1.2741 | 1.277 | 1.3 | 1.25 | 1.1586 | 1.2 | 1.1 | [10 empty] | 69
| 1 | 6 | 1.24 | 1.3853 | 1.2542 | 1.4 | 1.29 | 1.2679 | 1.3 | 1.1
| Mean |  | 1.3067 | 1.3472 | 1.2991 | 1.4 | 1.2533 | 1.2381 | 1.2833 | 1.1167
| Median |  | 1.315 | 1.3624 | 1.2723 | 1.4 | 1.25 | 1.2381 | 1.3 | 1.1
| Std Dev. |  | 0.052789 | 0.04542 | 0.1059 | 0.063246 | 0.020656 | 0.059435 | 0.075277 | 0.040825
| Rel.Std.Dev. |  | 4.04% | 3.371% | 8.152% | 4.518% | 1.648% | 4.8% | 5.866% | 3.656%
| PDM3 |  | 1.948% | 5.109% | 1.359% | 9.23% | -2.214% | -3.4% | 0.1271% | -12.88%
| Z-Score (Absolute) |  | 0.23 | 0.83 | 0.12 | 1.61 | 0.56 | 0.79 | 0.12 | 2.59
| Analytical results for Zn in OREAS 250c (Certified Value 75 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 75 | 75.519 | 78.8865 | 63.1579 | 76 | 76 | 72 | 67.8525 | 72 | 72 | 69.1396 | 79 | 75.7894 | 76.8 | 75 | 78 | 74 | 65 | 1
| 1 | 2 | 76 | 78.621 | 79.4631 | 63.0642 | 76 | 78 | 72 | 69.6489 | 72 | 74 | 68.4419 | 76 | 71.1391 | 76.5 | 75 | 79 | 75 | 66 | 35
| 1 | 3 | 76 | 80.271 | 78.6513 | 62.3481 | 76 | 76 | 72 | 62.5069 | 76 | 72 | 66.7303 | 78 | 75.6518 | 75.9 | 75 | 77 | 73 | 65 | 16
| 1 | 4 | 77 | 76.245 | 77.5665 | 64.7298 | 77 | 78 | 72 | 70.1427 | 75 | 73 | 69.4322 | 76 | 70.6728 | 76.4 | 77 | 79 | 74 | 67 | 74.9419
| 1 | 5 | 74 | 88.411 | 77.9977 | 63.6528 | 76 | 76 | 73 | 68.0693 | 73 | 74 | 65.38 | 79 | 78.8616 | 73.5 | 77 | 79 | 77 | 67 | 70
| 1 | 6 | 72 | 79.787 | 78.8014 | 64.4981 | 76 | 77 | 73 | 70.1836 | 73 | 74 | 68.5135 | 78 | 76.7783 | 76 | 77 | 81 | 75 | 66
| Mean |  | 75 | 79.809 | 78.5611 | 63.5752 | 76.1667 | 76.8333 | 72.3333 | 68.0673 | 73.5 | 73.1667 | 67.9396 | 77.6667 | 74.8155 | 75.85 | 76 | 78.8333 | 74.6667 | 66
| Median |  | 75.5 | 79.204 | 78.7264 | 63.4053 | 76 | 76.5 | 72 | 68.8591 | 73 | 73.5 | 68.4777 | 78 | 75.7206 | 76.2 | 76 | 79 | 74.5 | 66
| Std Dev. |  | 1.7889 | 4.6201 | 0.67704 | 0.90905 | 0.40825 | 0.98319 | 0.5164 | 2.907 | 1.6432 | 0.98319 | 1.5661 | 1.3663 | 3.2419 | 1.1979 | 1.0954 | 1.3292 | 1.3663 | 0.89443
| Rel.Std.Dev. |  | 2.385% | 5.789% | 0.8618% | 1.43% | 0.536% | 1.28% | 0.7139% | 4.271% | 2.236% | 1.344% | 2.305% | 1.759% | 4.333% | 1.579% | 1.441% | 1.686% | 1.83% | 1.355%
| PDM3 |  | 0.07749% | 6.494% | 4.829% | -15.17% | 1.634% | 2.524% | -3.481% | -9.173% | -1.924% | -2.369% | -9.344% | 3.636% | -0.1687% | 1.212% | 1.412% | 5.193% | -0.3673% | -11.93%
| Z-Score (Absolute) |  | 0.03 | 1.47 | 1.09 | 3.4 | 0.38 | 0.58 | 0.77 | 2.05 | 0.42 | 0.52 | 2.09 | 0.83 | 0.03 | 0.28 | 0.33 | 1.18 | 0.07 | 2.67
| Analytical results for Zr in OREAS 250c (Certified Value 188 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 187.4 | 179.308 | 199.682 | 205.5042 | 186 | 17 | 181 | 172 | 188.5 | 192.9 | 165.5052 | 200 | 189.6067 | 199 | 194 | 193 | 188.5 | 178 | 1
| 1 | 2 | 193.5 | 182.4768 | 197.5989 | 194.7915 | 184 | 22 | 183 | 167 | 186.5 | 196.8 | 164.7397 | 193.5 | 183.0212 | 195 | 194.5 | 197 | 186.5 | 176 | 36
| 1 | 3 | 196 | 182.5632 | 193.0623 | 197.1556 | 187 | 35 | 182 | 168 | 195 | 191.4 | 165.1516 | 204 | 177.5433 | 195 | 194.5 | 194.5 | 185.5 | 176 | 16
| 1 | 4 | 193.8 | 182.8548 | 194.2617 | 222.9007 | 184 | 9 | 183 | 172 | 188.5 | 193.5 | 164.6419 | 188.5 | 173.2039 | 201 | 211 | 194.5 | 187.5 | 179 | 187.837
| 1 | 5 | 188.2 | 188.126 | 195.5105 | 199.1677 | 184 | 9 | 185 | 170 | 188.5 | 196.5 | 154.6701 | 194 | 194.4872 | 190 | 200 | 190.5 | 194 | 174 | 71
| 1 | 6 | 194.2 | 189 | 194.1029 | 202.6427 | 181 | 4 | 184 | 173 | 186.5 | 192.9 | 164.4683 | 202 | 188.1417 | 196 | 203 | 200 | 192.5 | 175
| Mean |  | 192.1833 | 184.0548 | 195.703 | 203.6937 | 184.3333 | 16 | 183 | 170.3333 | 188.9167 | 194 | 163.1961 | 197 | 184.334 | 196 | 199.5 | 194.9167 | 189.0833 | 176.3333
| Median |  | 193.65 | 182.709 | 194.8861 | 200.9052 | 184 | 13 | 183 | 171 | 188.5 | 193.2 | 164.6908 | 197 | 185.5814 | 195.5 | 197.25 | 194.5 | 188 | 176
| Std Dev. |  | 3.5136 | 3.7341 | 2.493 | 10.1565 | 2.0656 | 11.3137 | 1.4142 | 2.4221 | 3.1371 | 2.1689 | 4.1938 | 5.9414 | 7.9637 | 3.7947 | 6.7082 | 3.2774 | 3.412 | 1.8619
| Rel.Std.Dev. |  | 1.828% | 2.029% | 1.274% | 4.986% | 1.121% | 70.71% | 0.7728% | 1.422% | 1.661% | 1.118% | 2.57% | 3.016% | 4.32% | 1.936% | 3.363% | 1.681% | 1.804% | 1.056%
| PDM3 |  | 2.314% | -2.014% | 4.188% | 8.442% | -1.865% | -91.48% | -2.575% | -9.319% | 0.5748% | 3.281% | -13.12% | 4.878% | -1.865% | 4.346% | 6.209% | 3.769% | 0.6635% | -6.124%
| Z-Score (Absolute) |  | 0.34 | 0.53 | 0.71 | 1.56 | 0.5 | 18.37 | 0.64 | 1.98 | 0.01 | 0.53 | 2.74 | 0.85 | 0.5 | 0.74 | 1.11 | 0.63 | 0.01 | 1.34
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C634: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C652: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C670: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C688: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C706: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 250c (Certified Value 0.079 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.08 | < 0.2 | < 10 | 0.08 | 0.078 | 0.08 | < 0.5 | 0.07 | < 0.1 | < 0.3 | 0.07 | < 0.05 | 0.1 | 0.07 | 0.08 | 0.08 |  |  |  | 1
| 1 | 2 | 0.07 | < 0.2 | < 10 | 0.09 | 0.08 | 0.08 | < 0.5 | 0.08 | < 0.1 | < 0.3 | 0.06 | < 0.05 | 0.1 | 0.07 | 0.07 | 0.08 |  |  |  | 20
| 1 | 3 | 0.08 | < 0.2 | < 10 | 0.09 | 0.084 | 0.09 | < 0.5 | 0.08 | < 0.1 | < 0.3 | 0.07 | < 0.05 | 0.1 | 0.1 | 0.09 | 0.08 |  |  |  | 16
| 1 | 4 | 0.08 | < 0.2 | < 10 | 0.08 | 0.08 | 0.08 | < 0.5 | 0.08 | 0.1 | < 0.3 | 0.07 | < 0.05 | 0.1 | 0.08 | 0.08 | 0.08 |  |  |  | 0.078574
| 1 | 5 | 0.07 | < 0.2 | < 10 | 0.09 | 0.082 | 0.09 | < 0.5 | 0.08 | < 0.1 | < 0.3 | 0.07 | < 0.05 | 0.1 | 0.08 | 0.08 | 0.08 |  |  |  | 73
| 1 | 6 | 0.07 | < 0.2 | < 10 | 0.09 | 0.089 | 0.07 | < 0.5 | 0.07 | < 0.1 | < 0.3 | 0.07 | < 0.05 | 0.1 | 0.07 | 0.07 | 0.08
| Mean |  | 0.075 |  |  | 0.086667 | 0.082167 | 0.081667 |  | 0.076667 | 0.1 |  | 0.068333 |  | 0.1 | 0.078333 | 0.078333 | 0.08
| Median |  | 0.075 |  |  | 0.09 | 0.081 | 0.08 |  | 0.08 | 0.1 |  | 0.07 |  | 0.1 | 0.075 | 0.08 | 0.08
| Std Dev. |  | 0.0054772 |  |  | 0.005164 | 0.00392 | 0.0075277 |  | 0.005164 |  |  | 0.0040825 |  | 1.5202e-17 | 0.01169 | 0.0075277 | 0
| Rel.Std.Dev. |  | 7.303% |  |  | 5.958% | 4.771% | 9.218% |  | 6.736% |  |  | 5.974% |  | 1.52e-14% | 14.92% | 9.61% | 0%
| PDM3 |  | -4.549% |  |  | 10.3% | 4.572% | 3.936% |  | -2.428% | 27.27% |  | -13.03% |  | 27.27% | -0.3064% | -0.3064% | 1.815%
| Z-Score (Absolute) |  | 0.43 | 1.41 | 361.86 | 0.43 | 0.1 | 0.06 | 12.44 | 0.31 | 1.66 | 5.09 | 0.92 | 4.11 | 1.41 | 0.18 | 0.18 | 0.06
| Analytical results for Al in OREAS 250c (Certified Value 1.64 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.6107 | 2.0025 | 1.6777 | 1.45 | 1.72 | 1.83 | 1.56 | 1.72 | 1.51 | 1.67 | 1.47 | 1.6977 | 1.86 | 1.66 | 1.6 | 1.54 | 1.512 |  |  | 1
| 1 | 2 | 1.5764 | 2.0083 | 1.6397 | 1.52 | 1.75 | 1.81 | 1.51 | 1.73 | 1.63 | 1.63 | 1.49 | 1.8496 | 1.92 | 1.62 | 1.58 | 1.59 | 1.469 |  |  | #N/A
| 1 | 3 | 1.6457 | 2.0659 | 1.6511 | 1.65 | 1.71 | 1.79 | 1.54 | 1.72 | 1.61 | 1.55 | 1.48 | 1.8085 | 1.87 | 1.64 | 1.56 | 1.58 | 1.475 |  |  | 16
| 1 | 4 | 1.6903 | 1.9702 | 1.6783 | 1.62 | 1.76 | 1.74 | 1.61 | 1.77 | 1.55 | 1.54 | 1.45 | 1.7708 | 1.88 | 1.65 | 1.62 | 1.56 | 1.483 |  |  | 1.6421
| 1 | 5 | 1.647 | 2.0291 | 1.6513 | 1.67 | 1.71 | 1.84 | 1.57 | 1.73 | 1.65 | 1.66 | 1.48 | 1.8461 | 1.89 | 1.59 | 1.54 | 1.58 | 1.471 |  |  | 74
| 1 | 6 | 1.6216 | 1.9418 | 1.6628 | 1.68 | 1.73 | 1.88 | 1.52 | 1.78 | 1.56 | 1.41 | 1.46 | 1.6764 | 1.86 | 1.62 | 1.55 | 1.58 | 1.471
| Mean |  | 1.632 | 2.003 | 1.6601 | 1.5983 | 1.73 | 1.815 | 1.5517 | 1.7417 | 1.585 | 1.5767 | 1.4717 | 1.7749 | 1.88 | 1.63 | 1.575 | 1.5717 | 1.4802
| Median |  | 1.6337 | 2.0054 | 1.657 | 1.635 | 1.725 | 1.82 | 1.55 | 1.73 | 1.585 | 1.59 | 1.475 | 1.7897 | 1.875 | 1.63 | 1.57 | 1.58 | 1.473
| Std Dev. |  | 0.038639 | 0.043551 | 0.015633 | 0.092826 | 0.020976 | 0.047645 | 0.03656 | 0.026394 | 0.053572 | 0.098319 | 0.01472 | 0.074103 | 0.022804 | 0.025298 | 0.030822 | 0.018348 | 0.016376
| Rel.Std.Dev. |  | 2.368% | 2.174% | 0.9417% | 5.808% | 1.212% | 2.625% | 2.356% | 1.515% | 3.38% | 6.236% | 1% | 4.175% | 1.213% | 1.552% | 1.957% | 1.167% | 1.106%
| PDM3 |  | -0.6188% | 21.98% | 1.098% | -2.666% | 5.352% | 10.53% | -5.508% | 6.063% | -3.478% | -3.985% | -10.38% | 8.084% | 14.49% | -0.7375% | -4.087% | -4.29% | -9.862%
| Z-Score (Absolute) |  | 0.02 | 3.21 | 0.26 | 0.27 | 0.86 | 1.59 | 0.67 | 0.96 | 0.39 | 0.46 | 1.36 | 1.25 | 2.15 | 0 | 0.47 | 0.5 | 1.29
| Analytical results for As in OREAS 250c (Certified Value 43.4 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 46 | 44.3296 | < 100 | 41 | 42 | 41.4 | 45.8 | 42.5467 | 41.9 | 42.3 | 43.492 | 42.6 | 42.6222 | 43.2 | 45.5 | 45.2 | 43 | 37 |  | 1
| 1 | 2 | 43 | 43.6484 | < 100 | 44 | 43 | 42.5 | 45.3 | 42.46 | 44.7 | 41.4 | 43.434 | 42.1 | 41.8982 | 44.5 | 44.9 | 44.3 | 42.8 | 37 |  | 21
| 1 | 3 | 45 | 44.937 | < 100 | 42 | 43 | 41.9 | 45.1 | 42.5233 | 43.8 | 41.2 | 43.149 | 42.6 | 41.725 | 43.4 | 45.7 | 44.4 | 44.9 | 37 |  | 16
| 1 | 4 | 45 | 45.8457 | < 100 | 46 | 43 | 39.1 | 44.9 | 42.9533 | 42.4 | 41.7 | 42.993 | 42 | 41.7106 | 42.4 | 46 | 45.3 | 43.7 | 38 |  | 43.4324
| 1 | 5 | 45 | 44.6051 | < 100 | 43 | 42 | 41.2 | 45.3 | 42.38 | 42.8 | 42.6 | 43.544 | 42.7 | 42.6018 | 43.6 | 44.4 | 44.8 | 43.5 | 36 |  | 75
| 1 | 6 | 45 | 45.2543 | < 100 | 45 | 43 | 42.3 | 45.1 | 42.5633 | 43 | 39.2 | 43.369 | 42 | 42.326 | 44.1 | 44.1 | 42.9 | 43.7 | 37
| Mean |  | 44.8333 | 44.77 |  | 43.5 | 42.6667 | 41.4 | 45.25 | 42.5711 | 43.1 | 41.4 | 43.3302 | 42.3333 | 42.1473 | 43.5333 | 45.1 | 44.4833 | 43.6 | 37
| Median |  | 45 | 44.7711 |  | 43.5 | 43 | 41.65 | 45.2 | 42.535 | 42.9 | 41.55 | 43.4015 | 42.35 | 42.1121 | 43.5 | 45.2 | 44.6 | 43.6 | 37
| Std Dev. |  | 0.98319 | 0.7614 |  | 1.8708 | 0.5164 | 1.2329 | 0.30822 | 0.19895 | 1.008 | 1.2017 | 0.21481 | 0.33267 | 0.42311 | 0.73121 | 0.75631 | 0.87502 | 0.73756 | 0.63246
| Rel.Std.Dev. |  | 2.193% | 1.701% |  | 4.301% | 1.21% | 2.978% | 0.6812% | 0.4673% | 2.339% | 2.903% | 0.4957% | 0.7858% | 1.004% | 1.68% | 1.677% | 1.967% | 1.692% | 1.709%
| PDM3 |  | 3.226% | 3.08% |  | 0.1556% | -1.763% | -4.679% | 4.185% | -1.983% | -0.7654% | -4.679% | -0.2354% | -2.531% | -2.959% | 0.2324% | 3.84% | 2.42% | 0.3859% | -14.81%
| Z-Score (Absolute) |  | 0.81 | 0.78 | 3.78 | 0.05 | 0.43 | 1.16 | 1.05 | 0.48 | 0.18 | 1.16 | 0.05 | 0.62 | 0.73 | 0.07 | 0.97 | 0.61 | 0.11 | 3.68
| Analytical results for B in OREAS 250c (Certified Value < 10 ppm) | [20 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.09 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [9 empty] | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 1
|  |  | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 1
| 1 | 1 | < 10 | 2.6 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 1
| 1 | 2 | < 10 | 2.4 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | #N/A
| 1 | 3 | < 10 | 2.6 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 16
| 1 | 4 | < 10 | 2.9 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | < 10
| 1 | 5 | < 10 | 3.6 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | [9 empty] | 76
| 1 | 6 | < 10 | 2.8 | 2 | < 10 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10
| Mean |  |  | 2.8167 | 2
| Median |  |  | 2.7 | 2
| Std Dev. |  |  | 0.42151 | 0
| Rel.Std.Dev. |  |  | 14.96% | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 250c (Certified Value 95 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 100 | 112.4388 | 97.394 | 100.89 | 103 | 94 | 92.2 | 96 | 92.7 | 100 | 76 | 90 | 90.6582 | 100 | 100 | 90 | 86 |  |  | 1
| 1 | 2 | 96 | 110.0844 | 96.3843 | 100.959 | 107 | 96 | 97.3 | 94 | 93.64 | 100 | 92 | 90 | 89.4714 | 100 | 100 | 100 | 85 |  |  | 22
| 1 | 3 | 98 | 112.3092 | 97.6138 | 104.682 | 112 | 93 | 93.7 | 95 | 95.52 | 100 | 94 | 90 | 97.094 | 100 | 100 | 90 | 87 |  |  | 16
| 1 | 4 | 101 | 109.3932 | 97.5835 | 105.693 | 123 | 94 | 89.4 | 97 | 94.2633 | 100 | 84 | 90 | 94.834 | 100 | 100 | 90 | 88 |  |  | 94.5276
| 1 | 5 | 98 | 109.8792 | 96.8734 | 101.342 | 113 | 92 | 94.4 | 97 | 93.83 | 100 | 88 | 90 | 96.1666 | 100 | 100 | 100 | 87 |  |  | 77
| 1 | 6 | 100 | 106.5636 | 97.4246 | 105.007 | 112 | 93 | 97.7 | 97 | 93.37 | 100 | 86 | 90 | 91.7622 | 100 | 100 | 90 | 85
| Mean |  | 98.8333 | 110.1114 | 97.2123 | 103.0955 | 111.6667 | 93.6667 | 94.1167 | 96 | 93.8872 | 100 | 86.6667 | 90 | 93.3311 | 100 | 100 | 93.3333 | 86.3333
| Median |  | 99 | 109.9818 | 97.4093 | 103.012 | 112 | 93.5 | 94.05 | 96.5 | 93.735 | 100 | 87 | 90 | 93.2981 | 100 | 100 | 90 | 86.5
| Std Dev. |  | 1.8348 | 2.1635 | 0.48528 | 2.2548 | 6.7429 | 1.3663 | 3.1352 | 1.2649 | 0.95382 | 0 | 6.4083 | 0 | 3.1293 | 0 | 0 | 5.164 | 1.2111
| Rel.Std.Dev. |  | 1.857% | 1.965% | 0.4992% | 2.187% | 6.038% | 1.459% | 3.331% | 1.318% | 1.016% | 0% | 7.394% | 0% | 3.353% | 0% | 0% | 5.533% | 1.403%
| PDM3 |  | 4.555% | 16.49% | 2.84% | 9.064% | 18.13% | -0.9108% | -0.4347% | 1.558% | -0.6775% | 5.789% | -8.316% | -4.79% | -1.266% | 5.789% | 5.789% | -1.263% | -8.669%
| Z-Score (Absolute) |  | 0.86 | 3.42 | 0.49 | 1.83 | 3.78 | 0.32 | 0.21 | 0.21 | 0.27 | NA | 1.91 | NA | 0.39 | NA | NA | NA | 1.98
Results from laboratories 1.12, 1.15, 1.19 and 1.21 were removed due to their 10 ppm reading resolution. Results from laboratory 1.20 were removed due to their 100 ppm reading resolution.
| Analytical results for Be in OREAS 250c (Certified Value 1.08 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 1.14 | 1.1588 | < 10 | 1.2 | 1 | 1.2 | 1.01 | < 5 | 1.03 | 1.01 | 0.99322 | 1.1 | 1.11 | 1.07 | 1 | 1.19 |  |  |  | 1
| 1 | 2 | 1.03 | 1.1352 | < 10 | 1.3 | 1.1 | 1.1 | 1.01 | < 5 | 1.06 | 1.01 | 1.0273 | 1.1 | 1.08 | 1.06 | 1.03 | 1.17 |  |  |  | 23
| 1 | 3 | 1.13 | 1.1321 | < 10 | 1.3 | 1 | 1.1 | 0.99 | < 5 | 1.06 | 1.02 | 1.064 | 1.1 | 1.09 | 1.04 | 1.02 | 1.32 |  |  |  | 16
| 1 | 4 | 1.07 | 1.1552 | < 10 | 1.4 | 1 | 1.1 | 1.01 | < 5 | 1.05 | 0.98 | 1.0999 | 1.2 | 1.1 | 1.07 | 1.02 | 1.26 |  |  |  | 1.0776
| 1 | 5 | 1.13 | 1.1381 | < 10 | 1.3 | 1 | 1.1 | 0.99 | < 5 | 1.08 | 1.01 | 1.0691 | 1.2 | 1.07 | 1.05 | 1.02 | 1.26 |  |  |  | 78
| 1 | 6 | 1.09 | 1.1529 | < 10 | 1.3 | 1 | 1.1 | 0.98 | < 5 | 1.05 | 1 | 1.0152 | 1.1 | 1.08 | 1.05 | 1.02 | 1.28
| Mean |  | 1.0983 | 1.1454 |  | 1.3 | 1.0167 | 1.1167 | 0.99833 |  | 1.055 | 1.005 | 1.0448 | 1.1333 | 1.0883 | 1.0567 | 1.0183 | 1.2467
| Median |  | 1.11 | 1.1455 |  | 1.3 | 1 | 1.1 | 1 |  | 1.055 | 1.01 | 1.0457 | 1.1 | 1.085 | 1.055 | 1.02 | 1.26
| Std Dev. |  | 0.043089 | 0.011533 |  | 0.063246 | 0.040825 | 0.040825 | 0.013292 |  | 0.016432 | 0.013784 | 0.039609 | 0.05164 | 0.01472 | 0.012111 | 0.0098319 | 0.056451
| Rel.Std.Dev. |  | 3.923% | 1.007% |  | 4.865% | 4.016% | 3.656% | 1.331% |  | 1.558% | 1.372% | 3.791% | 4.556% | 1.352% | 1.146% | 0.9655% | 4.528%
| PDM3 |  | 1.924% | 6.289% |  | 20.64% | -5.655% | 3.625% | -7.356% |  | -2.098% | -6.738% | -3.044% | 5.172% | 0.9957% | -1.943% | -5.5% | 15.69%
| Z-Score (Absolute) |  | 0.05 | 0.55 | 41.47 | 2.19 | 0.81 | 0.25 | 1.01 | 14.93 | 0.41 | 0.94 | 0.52 | 0.42 | 0.05 | 0.39 | 0.8 | 1.63
| Analytical results for Bi in OREAS 250c (Certified Value 0.26 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.28 | 0.2815 | < 10 | < 0.3 | 0.29 | 0.26 | 0.22 | < 2 | 0.26 | 0.2 | 0.264 | 0.24 | 0.24904 | 0.27 | 0.26 | 0.27 | 0.24 | 0.35 |  | 1
| 1 | 2 | 0.27 | 0.24735 | < 10 | < 0.3 | 0.29 | 0.27 | 0.23 | < 2 | 0.26 | 0.3 | 0.27 | 0.24 | 0.27384 | 0.32 | 0.27 | 0.28 | 0.24 | 0.35 |  | 24
| 1 | 3 | 0.28 | 0.25951 | < 10 | < 0.3 | 0.3 | 0.26 | 0.23 | < 2 | 0.26 | 0.3 | 0.274 | 0.24 | 0.31191 | 0.28 | 0.28 | 0.28 | 0.25 | 0.38 |  | 16
| 1 | 4 | 0.28 | 0.26814 | < 10 | < 0.3 | 0.3 | 0.25 | 0.23 | < 2 | 0.25 | 0.3 | 0.261 | 0.24 | 0.25933 | 0.29 | 0.27 | 0.28 | 0.25 | 0.35 |  | 0.2647
| 1 | 5 | 0.28 | 0.26877 | < 10 | < 0.3 | 0.29 | 0.26 | 0.23 | < 2 | 0.25 | 0.3 | 0.272 | 0.24 | 0.29915 | 0.28 | 0.26 | 0.29 | 0.24 | 0.34 |  | 79
| 1 | 6 | 0.29 | 0.26477 | < 10 | < 0.3 | 0.3 | 0.28 | 0.22 | < 2 | 0.26 | 0.3 | 0.258 | 0.24 | 0.28424 | 0.28 | 0.26 | 0.25 | 0.23 | 0.36
| Mean |  | 0.28 | 0.26501 |  |  | 0.295 | 0.26333 | 0.22667 |  | 0.25667 | 0.28333 | 0.2665 | 0.24 | 0.27958 | 0.28667 | 0.26667 | 0.275 | 0.24167 | 0.355
| Median |  | 0.28 | 0.26645 |  |  | 0.295 | 0.26 | 0.23 |  | 0.26 | 0.3 | 0.267 | 0.24 | 0.27904 | 0.28 | 0.265 | 0.28 | 0.24 | 0.35
| Std Dev. |  | 0.0063246 | 0.011296 |  |  | 0.0054772 | 0.010328 | 0.005164 |  | 0.005164 | 0.040825 | 0.006442 | 0 | 0.023771 | 0.017512 | 0.008165 | 0.013784 | 0.0075277 | 0.013784
| Rel.Std.Dev. |  | 2.259% | 4.263% |  |  | 1.857% | 3.922% | 2.278% |  | 2.012% | 14.41% | 2.417% | 0% | 8.502% | 6.109% | 3.062% | 5.012% | 3.115% | 3.883%
| PDM3 |  | 5.78% | 0.1163% |  |  | 11.45% | -0.5161% | -14.37% |  | -3.035% | 7.04% | 0.6803% | -9.331% | 5.623% | 8.299% | 0.7432% | 3.891% | -8.701% | 34.11%
| Z-Score (Absolute) |  | 0.45 | 0.06 | 159.59 | 3.93 | 0.96 | 0.11 | 1.35 | 24.72 | 0.34 | NA | 0.01 | 0.9 | 0.44 | 0.67 | 0 | 0.28 | 0.84 | 2.98
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 250c (Certified Value 0.524 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 0.54 | 0.65737 | 0.5314 | 0.39957 | 1.13 | 0.53 | 0.67 | 0.51 | 0.53 | 0.5 | 0.52 | 0.42479 | 0.48 | 0.52571 | 0.62 | 0.56 | 0.52 | 0.49 | 0.447 | 1
| 1 | 2 | 0.53 | 0.65168 | 0.53482 | 0.42755 | 1.17 | 0.54 | 0.66 | 0.5 | 0.54 | 0.5 | 0.51 | 0.43098 | 0.49 | 0.59577 | 0.63 | 0.54 | 0.52 | 0.5 | 0.444 | #N/A
| 1 | 3 | 0.55 | 0.65779 | 0.5358 | 0.42161 | 1.2 | 0.52 | 0.66 | 0.51 | 0.56 | 0.5 | 0.5 | 0.43438 | 0.48 | 0.55221 | 0.62 | 0.55 | 0.51 | 0.5 | 0.45 | 16
| 1 | 4 | 0.56 | 0.65106 | 0.55489 | 0.42872 | 1.54 | 0.53 | 0.62 | 0.52 | 0.54 | 0.52 | 0.5 | 0.43038 | 0.47 | 0.55069 | 0.62 | 0.56 | 0.53 | 0.5 | 0.452 | 0.52434
| 1 | 5 | 0.55 | 0.65732 | 0.53816 | 0.40624 | 1.22 | 0.53 | 0.67 | 0.52 | 0.55 | 0.51 | 0.53 | 0.43528 | 0.48 | 0.59688 | 0.61 | 0.52 | 0.51 | 0.5 | 0.447 | 80
| 1 | 6 | 0.54 | 0.64593 | 0.53928 | 0.40693 | 1.19 | 0.53 | 0.71 | 0.5 | 0.54 | 0.53 | 0.47 | 0.42409 | 0.48 | 0.52008 | 0.62 | 0.53 | 0.5 | 0.5 | 0.448
| Mean |  | 0.545 | 0.65352 | 0.53906 | 0.41511 | 1.2417 | 0.53 | 0.665 | 0.51 | 0.54333 | 0.51 | 0.505 | 0.42998 | 0.48 | 0.55689 | 0.62 | 0.54333 | 0.515 | 0.49833 | 0.448
| Median |  | 0.545 | 0.6545 | 0.53698 | 0.41427 | 1.195 | 0.53 | 0.665 | 0.51 | 0.54 | 0.505 | 0.505 | 0.43068 | 0.48 | 0.55145 | 0.62 | 0.545 | 0.515 | 0.5 | 0.4475
| Std Dev. |  | 0.010488 | 0.0047877 | 0.0082333 | 0.012405 | 0.14932 | 0.0063246 | 0.02881 | 0.0089443 | 0.010328 | 0.012649 | 0.020736 | 0.0046982 | 0.0063246 | 0.033161 | 0.0063246 | 0.01633 | 0.010488 | 0.0040825 | 0.0027568
| Rel.Std.Dev. |  | 1.924% | 0.7326% | 1.527% | 2.988% | 12.03% | 1.193% | 4.332% | 1.754% | 1.901% | 2.48% | 4.106% | 1.093% | 1.318% | 5.955% | 1.02% | 3.006% | 2.037% | 0.8192% | 0.6154%
| PDM3 |  | 3.94% | 24.64% | 2.807% | -20.83% | 136.8% | 1.079% | 26.83% | -2.735% | 3.622% | -2.735% | -3.689% | -18% | -8.457% | 6.208% | 18.24% | 3.622% | -1.782% | -4.96% | -14.56%
| Z-Score (Absolute) |  | 0.38 | 3.1 | 0.23 | 2.88 | 17.85 | 0 | 3.39 | 0.5 | 0.33 | 0.5 | 0.63 | 2.51 | 1.25 | 0.67 | 2.26 | 0.33 | 0.38 | 0.79 | 2.06
| Analytical results for Cd in OREAS 250c (Certified Value 0.067 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.04 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.06 | 0.12 | < 0.5 | 0.05 | < 0.1 | 0.05 | 0.07713 | 0.08 | 0.04 | 0.06 | 0.04 |  |  |  | 1
| 1 | 2 | 0.03 | < 0.2 | < 10 | < 0.3 | 0.1 | 0.02 | 0.11 | < 0.5 | 0.05 | < 0.1 | 0.04 | 0.074162 | 0.08 | 0.05 | 0.06 | 0.04 |  |  |  | 25
| 1 | 3 | 0.04 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.06 | 0.11 | < 0.5 | 0.04 | < 0.1 | 0.04 | 0.075677 | 0.07 | 0.05 | 0.06 | 0.05 |  |  |  | 16
| 1 | 4 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.07 | 0.12 | < 0.5 | 0.04 | 0.1 | 0.05 | 0.075692 | 0.07 | 0.05 | 0.05 | 0.04 |  |  |  | 0.06726
| 1 | 5 | 0.06 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.07 | 0.11 | < 0.5 | 0.05 | < 0.1 | 0.04 | 0.074925 | 0.07 | 0.05 | 0.06 | 0.04 |  |  |  | 81
| 1 | 6 | 0.04 | < 0.2 | < 10 | < 0.3 | 0.1 | 0.08 | 0.11 | < 0.5 | 0.05 | < 0.1 | 0.05 | 0.077119 | 0.09 | 0.04 | 0.06 | 0.04
| Mean |  | 0.045 |  |  |  | 0.09 | 0.06 | 0.11333 |  | 0.046667 | 0.1 | 0.045 | 0.075784 | 0.076667 | 0.046667 | 0.058333 | 0.041667
| Median |  | 0.04 |  |  |  | 0.09 | 0.065 | 0.11 |  | 0.05 | 0.1 | 0.045 | 0.075684 | 0.075 | 0.05 | 0.06 | 0.04
| Std Dev. |  | 0.012247 |  |  |  | 0.0089443 | 0.020976 | 0.005164 |  | 0.005164 |  | 0.0054772 | 0.0011819 | 0.008165 | 0.005164 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 27.22% |  |  |  | 9.938% | 34.96% | 4.556% |  | 11.07% |  | 12.17% | 1.56% | 10.65% | 11.07% | 6.999% | 9.798%
| PDM3 |  | -33.1% |  |  |  | 33.81% | -10.79% | 68.5% |  | -30.62% | 48.68% | -33.1% | 12.67% | 13.99% | -30.62% | -13.27% | -38.05%
| Z-Score (Absolute) |  | 0.69 | 0.96 | 147.81 | 2.46 | 0.66 | 0.24 | 1.36 | 5.46 | 0.64 | 0.29 | 0.69 | 0.24 | 0.26 | 0.64 | 0.29 | 0.79
| Analytical results for Ce in OREAS 250c (Certified Value 40.9 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 42.196 | 39.9803 | 42.5 | 42.4 | 39.93 | 40.2 | 40.4 | 43.4133 | 37 | 41.9 | 44.4 | 39 | 39.09 | [6 empty] | 1
| 1 | 2 | 40.753 | 38.6172 | 43.3 | 43.5 | 39.57 | 40.9 | 40 | 43.5616 | 38.2 | 41.7 | 47.5 | 39.9 | 39.48 | [6 empty] | 26
| 1 | 3 | 42.552 | 39.5361 | 41.7 | 41.6 | 39.86 | 40.3 | 40.7 | 43.3389 | 37.8 | 42.3 | 43.3 | 40 | 39.08 | [6 empty] | 16
| 1 | 4 | 42.713 | 39.323 | 42.7 | 40.5 | 40.63 | 40.9 | 39.1 | 42.6912 | 36.2 | 43.1 | 43.9 | 40 | 39.35 | [6 empty] | 40.8783
| 1 | 5 | 41.818 | 39.0633 | 41.6 | 42.5 | 40.37 | 38.6 | 39.9 | 42.8094 | 37.2 | 41 | 43.3 | 40.1 | 38.55 | [6 empty] | 82
| 1 | 6 | 42.797 | 39.7138 | 41.6 | 43.9 | 39.88 | 41.5 | 39.9 | 42.3566 | 37 | 41.8 | 43.4 | 40.2 | 38.92
| Mean |  | 42.1382 | 39.3723 | 42.2333 | 42.4 | 40.04 | 40.4 | 40 | 43.0285 | 37.2333 | 41.9667 | 44.3 | 39.8667 | 39.0783
| Median |  | 42.374 | 39.4296 | 42.1 | 42.45 | 39.905 | 40.6 | 39.95 | 43.0741 | 37.1 | 41.85 | 43.65 | 40 | 39.085
| Std Dev. |  | 0.76945 | 0.48604 | 0.70899 | 1.2426 | 0.38678 | 1 | 0.54406 | 0.47787 | 0.69761 | 0.69761 | 1.626 | 0.43665 | 0.32848
| Rel.Std.Dev. |  | 1.826% | 1.234% | 1.679% | 2.931% | 0.966% | 2.475% | 1.36% | 1.111% | 1.874% | 1.662% | 3.671% | 1.095% | 0.8406%
| PDM3 |  | 3.082% | -3.684% | 3.315% | 3.723% | -2.051% | -1.17% | -2.148% | 5.26% | -8.917% | 2.663% | 8.371% | -2.475% | -4.403%
| Z-Score (Absolute) |  | 0.75 | 0.44 | 0.79 | 0.86 | 0.15 | 0 | 0.17 | 1.13 | 1.36 | 0.67 | 1.68 | 0.23 | 0.57
| Analytical results for Co in OREAS 250c (Certified Value 15.9 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 15.5 | 18.9108 | 16.0643 | 19.13 | 28 | 16.3 | 16.5 | 16 | 16.69 | 15 | 12.5 | 15.7 | 16.315 | 16.7 | 16.2 | 15.7 | 15.1 | 15.4 |  | 1
| 1 | 2 | 15.6 | 19.386 | 15.6839 | 19.103 | 18 | 16.7 | 16.7 | 16 | 16.81 | 14.5 | 15.5 | 15.7 | 16.6008 | 17.23 | 16 | 16.5 | 14.8 | 15.4 |  | 27
| 1 | 3 | 15.5 | 19.5048 | 16.3362 | 16.995 | 33 | 16.3 | 17.1 | 16 | 17.0367 | 14.4 | 15.9 | 16.1 | 16.515 | 16.83 | 16.2 | 15.3 | 15.6 | 15.1 |  | 16
| 1 | 4 | 15.9 | 18.6084 | 16.4829 | 18.899 | 17 | 16.2 | 15.7 | 17 | 16.8967 | 15 | 14.6 | 15.1 | 16.3367 | 16.69 | 16.2 | 15.6 | 15.3 | 15.1 |  | 15.9278
| 1 | 5 | 15.6 | 18.9108 | 16.2815 | 21.264 | 315 | 16.1 | 16.4 | 17 | 16.7967 | 14 | 15 | 15.8 | 16.3537 | 16.88 | 15.6 | 15.3 | 15.5 | 15.3 |  | 83
| 1 | 6 | 15.8 | 18.576 | 16.5232 | 18.436 | 93 | 16.1 | 17.2 | 16 | 16.8567 | 15.4 | 14.8 | 15.4 | 16.5893 | 16.94 | 15.8 | 15 | 15.4 | 15.4
| Mean |  | 15.65 | 18.9828 | 16.2287 | 18.9712 | 84 | 16.2833 | 16.6 | 16.3333 | 16.8478 | 14.7167 | 14.7167 | 15.6333 | 16.4517 | 16.8783 | 16 | 15.5667 | 15.2833 | 15.2833
| Median |  | 15.6 | 18.9108 | 16.3089 | 19.001 | 30.5 | 16.25 | 16.6 | 16 | 16.8333 | 14.75 | 14.9 | 15.7 | 16.4343 | 16.855 | 16.1 | 15.45 | 15.35 | 15.35
| Std Dev. |  | 0.16432 | 0.38758 | 0.3129 | 1.3782 | 116.6396 | 0.22286 | 0.54406 | 0.5164 | 0.11583 | 0.50761 | 1.1856 | 0.34448 | 0.1317 | 0.19854 | 0.25298 | 0.52026 | 0.29269 | 0.1472
| Rel.Std.Dev. |  | 1.05% | 2.042% | 1.928% | 7.264% | 138.9% | 1.369% | 3.277% | 3.162% | 0.6875% | 3.449% | 8.056% | 2.203% | 0.8005% | 1.176% | 1.581% | 3.342% | 1.915% | 0.9631%
| PDM3 |  | -1.744% | 19.18% | 1.889% | 19.11% | 427.4% | 2.232% | 4.22% | 2.546% | 5.776% | -7.604% | -7.604% | -1.849% | 3.29% | 5.968% | 0.4535% | -2.267% | -4.046% | -4.046%
| Z-Score (Absolute) |  | 0.66 | 2.95 | 0.03 | 2.94 | 73.38 | 0.03 | 0.37 | 0.08 | 0.64 | 1.67 | 1.67 | 0.67 | 0.21 | 0.67 | 0.28 | 0.75 | 1.05 | 1.05
| Analytical results for Cr in OREAS 250c (Certified Value 45.9 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 1
| 1 | 1 | 47 | 53.784 | 45.8377 | 33.519 | 41 | 57 | 52 | 45 | 45.9451 | 43 | 44 | 41 | 53.9731 | 55 | 47 | 44 | 42 | 41 |  | 1
| 1 | 2 | 47 | 55.4256 | 45.4177 | 34.611 | 42 | 57 | 52 | 43 | 46.7993 | 43 | 43 | 42 | 55.3818 | 57 | 46 | 45 | 43 | 41 |  | 28
| 1 | 3 | 49 | 56.7756 | 47.0111 | 35.412 | 44 | 58 | 52 | 44 | 47.6322 | 43 | 42 | 41 | 53.1046 | 56 | 47 | 44 | 43 | 41 |  | 16
| 1 | 4 | 50 | 54.3672 | 47.2332 | 35.725 | 44 | 57 | 48 | 46 | 46.1689 | 45 | 42 | 40 | 49.8365 | 54 | 46 | 45 | 43 | 40 |  | 45.8997
| 1 | 5 | 48 | 56.1924 | 46.4955 | 34.841 | 46 | 58 | 50 | 45 | 46.1851 | 44 | 44 | 42 | 54.3727 | 56 | 44 | 44 | 43 | 42 |  | 84
| 1 | 6 | 48 | 53.028 | 46.8188 | 34.659 | 46 | 57 | 53 | 43 | 46.2668 | 45 | 39 | 41 | 50.506 | 56 | 45 | 43 | 43 | 41
| Mean |  | 48.1667 | 54.9288 | 46.469 | 34.7945 | 43.8333 | 57.3333 | 51.1667 | 44.3333 | 46.4996 | 43.8333 | 42.3333 | 41.1667 | 52.8624 | 55.6667 | 45.8333 | 44.1667 | 42.8333 | 41
| Median |  | 48 | 54.8964 | 46.6572 | 34.75 | 44 | 57 | 52 | 44.5 | 46.2259 | 43.5 | 42.5 | 41 | 53.5389 | 56 | 46 | 44 | 43 | 41
| Std Dev. |  | 1.169 | 1.4487 | 0.70779 | 0.76528 | 2.0412 | 0.5164 | 1.8348 | 1.2111 | 0.62331 | 0.98319 | 1.8619 | 0.75277 | 2.2194 | 1.0328 | 1.169 | 0.75277 | 0.40825 | 0.63246
| Rel.Std.Dev. |  | 2.427% | 2.637% | 1.523% | 2.199% | 4.657% | 0.9007% | 3.586% | 2.732% | 1.34% | 2.243% | 4.398% | 1.829% | 4.199% | 1.855% | 2.551% | 1.704% | 0.9531% | 1.543%
| PDM3 |  | 4.939% | 19.67% | 1.24% | -24.19% | -4.502% | 24.91% | 11.48% | -3.413% | 1.307% | -4.502% | -7.77% | -10.31% | 15.17% | 21.28% | -0.1445% | -3.776% | -6.681% | -10.67%
| Z-Score (Absolute) |  | 0.71 | 2.28 | 0.32 | 2.38 | 0.29 | 2.83 | 1.41 | 0.17 | 0.33 | 0.29 | 0.64 | 0.91 | 1.8 | 2.45 | 0.17 | 0.21 | 0.52 | 0.94
| Analytical results for Cs in OREAS 250c (Certified Value 1.13 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 1.27 | 1.1312 | 1.28 | 1.06 | 1.18 | 1.06 | 1 | 1.1433 | 1.41 | 1.08 | 1.17 | 1.05 | 1.08 | [6 empty] | 1
| 1 | 2 | 1.16 | 1.0488 | 1.39 | 1.06 | 1.16 | 1.06 | 1 | 1.2255 | 1.45 | 1.06 | 1.14 | 1.02 | 1.08 | [6 empty] | 29
| 1 | 3 | 1.21 | 1.1469 | 1.35 | 1.09 | 1.15 | 1.04 | 1.02 | 1.2282 | 1.41 | 1.09 | 1.12 | 1.07 | 1.08 | [6 empty] | 16
| 1 | 4 | 1.25 | 1.161 | 1.37 | 1.03 | 1.2 | 1.06 | 0.98 | 1.2126 | 1.4 | 1.07 | 1.13 | 1.06 | 1.1 | [6 empty] | 1.1273
| 1 | 5 | 1.21 | 1.1055 | 1.26 | 1.11 | 1.17 | 1.02 | 1.02 | 1.1872 | 1.45 | 1.05 | 1.16 | 1.05 | 1.06 | [6 empty] | 85
| 1 | 6 | 1.25 | 1.1393 | 1.34 | 1.09 | 1.16 | 1.08 | 1.01 | 1.1565 | 1.41 | 1.07 | 1.14 | 1.04 | 1.07
| Mean |  | 1.225 | 1.1221 | 1.3317 | 1.0733 | 1.17 | 1.0533 | 1.005 | 1.1922 | 1.4217 | 1.07 | 1.1433 | 1.0483 | 1.0783
| Median |  | 1.23 | 1.1352 | 1.345 | 1.075 | 1.165 | 1.06 | 1.005 | 1.1999 | 1.41 | 1.07 | 1.14 | 1.05 | 1.08
| Std Dev. |  | 0.039875 | 0.040397 | 0.051153 | 0.028752 | 0.017889 | 0.020656 | 0.015166 | 0.036071 | 0.022286 | 0.014142 | 0.018619 | 0.017224 | 0.013292
| Rel.Std.Dev. |  | 3.255% | 3.6% | 3.841% | 2.679% | 1.529% | 1.961% | 1.509% | 3.026% | 1.568% | 1.322% | 1.628% | 1.643% | 1.233%
| PDM3 |  | 8.669% | -0.4587% | 18.13% | -4.785% | 3.79% | -6.56% | -10.85% | 5.762% | 26.12% | -5.081% | 1.424% | -7.003% | -4.342%
| Z-Score (Absolute) |  | 1.01 | 0 | 2.05 | 0.48 | 0.47 | 0.67 | 1.15 | 0.69 | 2.94 | 0.51 | 0.21 | 0.72 | 0.43
| Analytical results for Cu in OREAS 250c (Certified Value 18.9 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 19.8 | 16.4592 | 19.2417 | 16.131 | 21 | 22 | 21.4 | 20 | 20.13 | 17.2 | 14.2 | 16.7421 | 17.3 | 19.3446 | 20.7 | 20.9 | 21.1 | 18.6 | 15.1 | 1
| 1 | 2 | 18.9 | 14.796 | 19.18 | 15.67 | 22 | 21 | 20.6 | 19 | 20.0967 | 18 | 16.9 | 17.1363 | 17.1 | 19.3567 | 21.3 | 20.8 | 21.5 | 18.6 | 15.5 | 30
| 1 | 3 | 19.6 | 15.1956 | 19.5739 | 15.949 | 23 | 22 | 20.7 | 20 | 20.5767 | 17.8 | 17.9 | 17.1459 | 17.4 | 17.9117 | 20.4 | 21.1 | 20.6 | 19.2 | 15.3 | 16
| 1 | 4 | 20 | 14.9148 | 19.8176 | 16.438 | 23 | 22 | 20.6 | 20 | 20.3633 | 17.6 | 16.6 | 16.8394 | 16.8 | 18.3388 | 20.5 | 21 | 21.2 | 19.2 | 15.1 | 18.8888
| 1 | 5 | 19.4 | 14.9256 | 19.355 | 16.001 | 30 | 22 | 20.9 | 20 | 19.9167 | 17.2 | 16.6 | 16.9254 | 17.4 | 17.4218 | 21.6 | 20.2 | 21 | 19.1 | 15.3 | 86
| 1 | 6 | 19.8 | 13.9968 | 19.7218 | 16.999 | 30 | 22 | 21 | 20 | 20.38 | 18.1 | 16.5 | 16.5592 | 16.9 | 17.2904 | 20.8 | 20.5 | 19.8 | 19.4 | 15.1
| Mean |  | 19.5833 | 15.048 | 19.4817 | 16.198 | 24.8333 | 21.8333 | 20.8667 | 19.8333 | 20.2439 | 17.65 | 16.45 | 16.8914 | 17.15 | 18.2773 | 20.8833 | 20.75 | 20.8667 | 19.0167 | 15.2333
| Median |  | 19.7 | 14.9202 | 19.4645 | 16.066 | 23 | 22 | 20.8 | 20 | 20.2467 | 17.7 | 16.6 | 16.8824 | 17.2 | 18.1252 | 20.75 | 20.85 | 21.05 | 19.15 | 15.2
| Std Dev. |  | 0.392 | 0.80181 | 0.26216 | 0.46546 | 4.0702 | 0.40825 | 0.30768 | 0.40825 | 0.23885 | 0.38859 | 1.2178 | 0.22859 | 0.25884 | 0.91078 | 0.47081 | 0.33912 | 0.59889 | 0.33714 | 0.1633
| Rel.Std.Dev. |  | 2.002% | 5.328% | 1.346% | 2.874% | 16.39% | 1.87% | 1.475% | 2.058% | 1.18% | 2.202% | 7.403% | 1.353% | 1.509% | 4.983% | 2.255% | 1.634% | 2.87% | 1.773% | 1.072%
| PDM3 |  | 3.677% | -20.33% | 3.138% | -14.25% | 31.47% | 15.59% | 10.47% | 5% | 7.174% | -6.559% | -12.91% | -10.57% | -9.206% | -3.237% | 10.56% | 9.853% | 10.47% | 0.6767% | -19.35%
| Z-Score (Absolute) |  | 0.05 | 2.13 | 0 | 1.58 | 2.57 | 1.13 | 0.67 | 0.17 | 0.37 | 0.88 | 1.46 | 1.25 | 1.12 | 0.58 | 0.67 | 0.61 | 0.67 | 0.22 | 2.04
| Analytical results for Dy in OREAS 250c (Indicative Value 2.24 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 2.324 | 2.1913 | 2.3 | [16 empty] | 1
| 1 | 2 | 2.3 | 2.0733 | 2.2 | [16 empty] | 4
| 1 | 3 | 2.321 | 2.1255 | 2.3 | [16 empty] | 16
| 1 | 4 | 2.369 | 2.1445 | 2.1 | [16 empty] | 2.2407
| 1 | 5 | 2.333 | 2.1113 | 2.3 | [16 empty] | 10
| 1 | 6 | 2.306 | 2.2329 | 2.3
| Mean |  | 2.3255 | 2.1465 | 2.25
| Median |  | 2.3225 | 2.135 | 2.3
| Std Dev. |  | 0.024501 | 0.057461 | 0.083666
| Rel.Std.Dev. |  | 1.054% | 2.677% | 3.718%
| PDM3 |  | 3.787% | -4.204% | 0.4172%
| Z-Score (Absolute) |  | 0.67 | 0.92 | 0
| Analytical results for Er in OREAS 250c (Indicative Value 0.87 ppm) | [20 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.935 | 0.84203 | 0.8 | [16 empty] | 1
| 1 | 2 | 0.904 | 0.82468 | 0.8 | [16 empty] | 5
| 1 | 3 | 0.931 | 0.87839 | 0.8 | [16 empty] | 16
| 1 | 4 | 0.928 | 0.85104 | 0.8 | [16 empty] | 0.86776
| 1 | 5 | 0.955 | 0.85436 | 0.8 | [16 empty] | 11
| 1 | 6 | 0.953 | 0.86314 | 0.9
| Mean |  | 0.93433 | 0.85228 | 0.81667
| Median |  | 0.933 | 0.8527 | 0.8
| Std Dev. |  | 0.018694 | 0.018287 | 0.040825
| Rel.Std.Dev. |  | 2.001% | 2.146% | 4.999%
| PDM3 |  | 7.672% | -1.784% | -5.888%
| Z-Score (Absolute) |  | 1.55 | 0 | 0.67
| Analytical results for Eu in OREAS 250c (Indicative Value 0.97 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 1.059 | 0.93643 | 0.9 | [16 empty] | 1
| 1 | 2 | 1.041 | 0.87409 | 0.9 | [16 empty] | 6
| 1 | 3 | 1.053 | 0.96282 | 0.9 | [16 empty] | 16
| 1 | 4 | 1.086 | 0.95282 | 0.9 | [16 empty] | 0.96937
| 1 | 5 | 1.06 | 0.92503 | 0.9 | [16 empty] | 12
| 1 | 6 | 1.067 | 0.93143 | 1
| Mean |  | 1.061 | 0.93044 | 0.91667
| Median |  | 1.0595 | 0.93393 | 0.9
| Std Dev. |  | 0.015033 | 0.030955 | 0.040825
| Rel.Std.Dev. |  | 1.417% | 3.327% | 4.454%
| PDM3 |  | 9.453% | -4.016% | -5.437%
| Z-Score (Absolute) |  | 6.39 | 0 | 0.67
| Analytical results for Fe in OREAS 250c (Certified Value 3.01 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 3.02 | 3.4411 | 3.0488 | 2.4864 | 2.97 | 3.28 | 3.51 | 2.8 | 2.95 | 2.85 | 3.03 | 2.6677 | 2.83 | 3.2262 | 3.33 | 3.07 | 2.82 | 2.81 | 2.74 | 1
| 1 | 2 | 2.96 | 3.436 | 2.999 | 2.5842 | 3.12 | 3.4 | 3.4 | 2.8 | 2.95 | 2.99 | 2.94 | 2.7435 | 2.8 | 3.4344 | 3.38 | 2.98 | 2.79 | 2.86 | 2.66 | #N/A
| 1 | 3 | 3.09 | 3.4653 | 3.0623 | 2.5984 | 3.24 | 3.29 | 3.4 | 2.8 | 2.98 | 2.99 | 2.88 | 2.7413 | 2.83 | 3.3722 | 3.3 | 2.99 | 2.76 | 2.85 | 2.74 | 16
| 1 | 4 | 3.14 | 3.3802 | 3.1116 | 2.6112 | 3.2 | 3.33 | 3.41 | 2.8 | 2.94 | 2.91 | 2.89 | 2.6872 | 2.73 | 3.3402 | 3.37 | 3.04 | 2.84 | 2.83 | 2.69 | 3.0079
| 1 | 5 | 3.08 | 3.4302 | 3.0545 | 2.5133 | 3.31 | 3.27 | 3.48 | 2.81 | 2.95 | 3.02 | 3.01 | 2.7223 | 2.79 | 3.3216 | 3.32 | 2.95 | 2.85 | 2.87 | 2.7 | 87
| 1 | 6 | 3.02 | 3.3197 | 3.0854 | 2.5177 | 3.34 | 3.28 | 3.47 | 2.79 | 2.93 | 2.93 | 2.71 | 2.5992 | 2.8 | 3.2255 | 3.36 | 2.96 | 2.85 | 2.84 | 2.68
| Mean |  | 3.0517 | 3.4121 | 3.0603 | 2.5519 | 3.1967 | 3.3083 | 3.445 | 2.8 | 2.95 | 2.9483 | 2.91 | 2.6935 | 2.7967 | 3.32 | 3.3433 | 2.9983 | 2.8183 | 2.8433 | 2.7017
| Median |  | 3.05 | 3.4331 | 3.0584 | 2.551 | 3.22 | 3.285 | 3.44 | 2.8 | 2.95 | 2.96 | 2.915 | 2.7048 | 2.8 | 3.3309 | 3.345 | 2.985 | 2.83 | 2.845 | 2.695
| Std Dev. |  | 0.064005 | 0.05315 | 0.037912 | 0.052301 | 0.13604 | 0.049565 | 0.047645 | 0.0063246 | 0.016733 | 0.063377 | 0.11541 | 0.055152 | 0.036697 | 0.08243 | 0.031411 | 0.047081 | 0.03656 | 0.021602 | 0.032506
| Rel.Std.Dev. |  | 2.097% | 1.558% | 1.239% | 2.05% | 4.256% | 1.498% | 1.383% | 0.2259% | 0.5672% | 2.15% | 3.966% | 2.048% | 1.312% | 2.483% | 0.9395% | 1.57% | 1.297% | 0.7598% | 1.203%
| PDM3 |  | 1.456% | 13.44% | 1.742% | -15.16% | 6.277% | 9.989% | 14.53% | -6.911% | -1.924% | -1.979% | -3.254% | -10.45% | -7.022% | 10.38% | 11.15% | -0.3172% | -6.301% | -5.47% | -10.18%
| Z-Score (Absolute) |  | 0.45 | 2.03 | 0.48 | 1.75 | 1.08 | 1.58 | 2.18 | 0.66 | 0 | 0.01 | 0.18 | 1.13 | 0.67 | 1.63 | 1.73 | 0.21 | 0.58 | 0.47 | 1.09
| Analytical results for Ga in OREAS 250c (Certified Value 5.97 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 5.83 | 5.8263 | 6.8 | 6.48 | 6.24 | 5.7623 | 5.61 | 4 | 5.16 | 5.8883 | 6.67 | 5.75 | 6.33 | 5.12 | [5 empty] | 1
| 1 | 2 | 5.82 | 5.6807 | 7.1 | 6.46 | 6.1 | 5.8523 | 5.26 | 5 | 5.09 | 6.2991 | 6.99 | 5.67 | 6.6 | 5.17 | [5 empty] | 34
| 1 | 3 | 5.88 | 5.7682 | 6.6 | 6.4 | 6.19 | 5.9011 | 5.18 | 5 | 5.24 | 6.4022 | 6.69 | 5.84 | 6.26 | 5.33 | [5 empty] | 16
| 1 | 4 | 5.86 | 5.998 | 7 | 6.11 | 6.41 | 5.8038 | 5.63 | 5 | 5.07 | 6.0703 | 6.35 | 6.23 | 6.31 | 5.25 | [5 empty] | 5.9673
| 1 | 5 | 5.88 | 5.8414 | 6.6 | 6.55 | 6.28 | 5.8639 | 5.04 | 5 | 5.22 | 6.282 | 6.56 | 5.51 | 6.11 | 5.3 | [5 empty] | 88
| 1 | 6 | 5.86 | 5.8312 | 6.7 | 6.71 | 6.15 | 5.8532 | 5.79 | 5 | 5.09 | 5.8164 | 6.37 | 5.6 | 6.06 | 5.25
| Mean |  | 5.855 | 5.8243 | 6.8 | 6.4517 | 6.2283 | 5.8395 | 5.4183 | 4.8333 | 5.145 | 6.1264 | 6.605 | 5.7667 | 6.2783 | 5.2367
| Median |  | 5.86 | 5.8288 | 6.75 | 6.47 | 6.215 | 5.8528 | 5.435 | 5 | 5.125 | 6.1761 | 6.615 | 5.71 | 6.285 | 5.25
| Std Dev. |  | 0.0251 | 0.10418 | 0.20976 | 0.19833 | 0.10944 | 0.048949 | 0.29822 | 0.40825 | 0.07287 | 0.23921 | 0.23747 | 0.2543 | 0.19156 | 0.078909
| Rel.Std.Dev. |  | 0.4287% | 1.789% | 3.085% | 3.074% | 1.757% | 0.8382% | 5.504% | 8.447% | 1.416% | 3.905% | 3.595% | 4.41% | 3.051% | 1.507%
| PDM3 |  | -1.882% | -2.397% | 13.95% | 8.117% | 4.374% | -2.143% | -9.2% | -19% | -13.78% | 2.666% | 10.69% | -3.363% | 5.212% | -12.24%
| Z-Score (Absolute) |  | 0 | 0.05 | 1.51 | 0.95 | 0.59 | 0.02 | 0.7 | NA | 1.13 | 0.43 | 1.19 | 0.14 | 0.67 | 0.98
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 250c (Indicative Value 3.38 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 3.377 | 3.4662 | 3.3 | [16 empty] | 1
| 1 | 2 | 3.344 | 3.2637 | 3.4 | [16 empty] | 7
| 1 | 3 | 3.412 | 3.4469 | 3.3 | [16 empty] | 16
| 1 | 4 | 3.477 | 3.4674 | 3.1 | [16 empty] | 3.3839
| 1 | 5 | 3.432 | 3.4342 | 3.3 | [16 empty] | 13
| 1 | 6 | 3.578 | 3.4109 | 3.4
| Mean |  | 3.4367 | 3.4149 | 3.3
| Median |  | 3.422 | 3.4406 | 3.3
| Std Dev. |  | 0.082921 | 0.077024 | 0.10954
| Rel.Std.Dev. |  | 2.413% | 2.256% | 3.32%
| PDM3 |  | 1.561% | 0.9173% | -2.478%
| Z-Score (Absolute) |  | 0.67 | 0 | 3.56
| Analytical results for Ge in OREAS 250c (Certified Value 0.1 ppm) | [20 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.09 | 1.10 | 1.12 | 1.15 | 1.19 | 1.20 | 1.21 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [11 empty] | 3
|  |  | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.2 | < 0.1 | 0.11 | 0.14 | 0.11 | 0.1 | 0.18 | 0.09 | [11 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.12 | 0.07 | 0.1 | 0.09 | 0.18 | 0.08 | [11 empty] | 1
| 1 | 3 | < 0.2 | < 0.1 | 0.11 | 0.08 | 0.1 | 0.09 | 0.19 | 0.09 | [11 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.1 | 0.12 | 0.09 | 0.09 | 0.17 | 0.1 | [11 empty] | 0.099667
| 1 | 5 | < 0.2 | < 0.1 | 0.11 | 0.07 | 0.1 | 0.09 | 0.17 | 0.1 | [11 empty] | 89
| 1 | 6 | < 0.2 | < 0.1 | 0.12 | 0.13 | 0.12 | 0.09 | 0.16 | 0.08
| Mean |  |  |  | 0.11167 | 0.10167 | 0.10333 | 0.091667 | 0.175 | 0.09
| Median |  |  |  | 0.11 | 0.1 | 0.1 | 0.09 | 0.175 | 0.09
| Std Dev. |  |  |  | 0.0075277 | 0.031885 | 0.010328 | 0.0040825 | 0.010488 | 0.0089443
| Rel.Std.Dev. |  |  |  | 6.741% | 31.36% | 9.995% | 4.454% | 5.993% | 9.938%
| PDM3 |  |  |  | 12.04% | 2.007% | 3.679% | -8.027% | 75.59% | -9.699%
| Z-Score (Absolute) |  | 0.06 | 3.43 | 0.73 | 0.06 | 0.17 | 0.62 | 5 | 0.73
| Analytical results for Hf in OREAS 250c (Indicative Value 0.68 ppm) | [20 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 0.26 | 0.96212 | 1.2 | < 0.1 | 0.8 | 0.57 | 0.6 | 0.89159 | 0.87 | 0.38 | 0.5 | 0.37 | 0.72 | [6 empty] | 1
| 1 | 2 | 0.24 | 0.84847 | 1.27 | < 0.1 | 0.83 | 0.46 | 0.45 | 1.0378 | 0.9 | 0.35 | 0.53 | 0.3 | 0.76 | [6 empty] | 8
| 1 | 3 | 0.27 | 0.91862 | 1.22 | < 0.1 | 0.72 | 0.67 | 0.64 | 1.0367 | 0.88 | 0.56 | 0.53 | 0.31 | 0.75 | [6 empty] | 16
| 1 | 4 | 0.23 | 0.89984 | 1.18 | < 0.1 | 0.78 | 0.54 | 0.54 | 1.0316 | 0.9 | 0.27 | 0.54 | 0.4 | 0.76 | [6 empty] | 0.67802
| 1 | 5 | 0.19 | 0.95782 | 1.23 | < 0.1 | 0.72 | 0.46 | 0.57 | 0.86885 | 0.84 | 0.48 | 0.67 | 0.3 | 0.74 | [6 empty] | 14
| 1 | 6 | 0.23 | 0.98323 | 1.19 | < 0.1 | 0.71 | 0.6 | 0.64 | 0.97889 | 0.86 | 0.45 | 0.61 | 0.26 | 0.73
| Mean |  | 0.23667 | 0.92835 | 1.215 |  | 0.76 | 0.55 | 0.57333 | 0.97423 | 0.875 | 0.415 | 0.56333 | 0.32333 | 0.74333
| Median |  | 0.235 | 0.93822 | 1.21 |  | 0.75 | 0.555 | 0.585 | 1.0053 | 0.875 | 0.415 | 0.535 | 0.305 | 0.745
| Std Dev. |  | 0.028048 | 0.049617 | 0.032711 |  | 0.0502 | 0.081976 | 0.072019 | 0.076402 | 0.023452 | 0.10291 | 0.06377 | 0.05164 | 0.01633
| Rel.Std.Dev. |  | 11.85% | 5.345% | 2.692% |  | 6.605% | 14.9% | 12.56% | 7.842% | 2.68% | 24.8% | 11.32% | 15.97% | 2.197%
| PDM3 |  | -65.09% | 36.92% | 79.2% |  | 12.09% | -18.88% | -15.44% | 43.69% | 29.05% | -38.79% | -16.92% | -52.31% | 9.633%
| Z-Score (Absolute) |  | 0.91 | 0.96 | 1.73 | 1.41 | 0.5 | 0.06 | 0 | 1.08 | 0.81 | 0.43 | 0.03 | 0.67 | 0.46
| Analytical results for Hg in OREAS 250c (Indicative Value 0.022 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [5 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | < 0.1 | < 1 | 0.01 | < 2 | 0.02 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 1 | 0.01 | 0.04 | 0.02 | 0.27 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 1 | 0.03 | < 2 | 0.02 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 1 | 0.01 | 0.03 | 0.02 | 0.29 | [5 empty] | 2
| 1 | 3 | < 0.1 | < 1 | < 0.01 | < 2 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 1 | 0.01 | 0.05 | 0.03 | 0.25 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 1 | < 0.01 | < 2 | 0.02 | 0.04 | < 0.01 | 0.02 | < 0.05 | 1 | 0.01 | 0.04 | 0.02 | 0.28 | [5 empty] | 0.021667
| 1 | 5 | < 0.1 | < 1 | < 0.01 | < 2 | 0.02 | 0.02 | < 0.01 | 0.01 | < 0.05 | 1 | 0.01 | 0.01 | 0.03 | 0.26 | [5 empty] | 15
| 1 | 6 | < 0.1 | < 1 | < 0.01 | < 2 | 0.02 | < 0.01 | < 0.01 | 0.01 | < 0.05 | 1 | 0.01 | 0.04 | 0.01 | 0.25
| Mean |  |  |  | 0.02 |  | 0.018333 | 0.03 |  | 0.011667 |  | 1 | 0.01 | 0.035 | 0.021667 | 0.26667
| Median |  |  |  | 0.02 |  | 0.02 | 0.03 |  | 0.01 |  | 1 | 0.01 | 0.04 | 0.02 | 0.265
| Std Dev. |  |  |  | 0.014142 |  | 0.0040825 | 0.014142 |  | 0.0040825 |  | 0 | 0 | 0.013784 | 0.0075277 | 0.01633
| Rel.Std.Dev. |  |  |  | 70.71% |  | 22.27% | 47.14% |  | 34.99% |  | 0% | 0% | 39.38% | 34.74% | 6.124%
| PDM3 |  |  |  | -7.692% |  | -15.38% | 38.46% |  | -46.15% |  | 4515% | -53.85% | 61.54% | -1.443e-13% | 1131%
| Z-Score (Absolute) |  | 1.35 | 24.11 | 0.67 | 49.39 | 0.25 | 0.51 | 0.93 | 0.59 | 0.08 | 49.39 | 0.67 | 0.59 | 0.08 | 12.31
| Analytical results for Ho in OREAS 250c (Indicative Value 0.36 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.375 | 0.35278 | 0.3 | [16 empty] | 1
| 1 | 2 | 0.36 | 0.32463 | 0.3 | [16 empty] | 4
| 1 | 3 | 0.365 | 0.35435 | 0.4 | [16 empty] | 16
| 1 | 4 | 0.375 | 0.35338 | 0.3 | [16 empty] | 0.35544
| 1 | 5 | 0.37 | 0.3468 | 0.4 | [16 empty] | 10
| 1 | 6 | 0.382 | 0.33889 | 0.4
| Mean |  | 0.37117 | 0.34514 | 0.35
| Median |  | 0.3725 | 0.34979 | 0.35
| Std Dev. |  | 0.0078846 | 0.011606 | 0.054772
| Rel.Std.Dev. |  | 2.124% | 3.363% | 15.65%
| PDM3 |  | 4.426% | -2.897% | -1.529%
| Z-Score (Absolute) |  | 2.94 | 0.67 | 0
| Analytical results for In in OREAS 250c (Certified Value 0.025 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.03 | < 0.05 | 0.04 | < 5 | 0.022 | 0.024 | < 0.05 | 0.1 | 0.023 | 0.032 | 0.024 | 0.024 | [7 empty] | 1
| 1 | 2 | 0.02 | < 0.05 | 0.03 | < 5 | 0.023 | 0.025 | < 0.05 | 0.1 | 0.023 | 0.03 | 0.023 | 0.029 | [7 empty] | 22
| 1 | 3 | 0.02 | < 0.05 | 0.03 | < 5 | 0.023 | 0.022 | < 0.05 | 0.1 | 0.026 | 0.03 | 0.025 | 0.024 | [7 empty] | 16
| 1 | 4 | 0.03 | < 0.05 | 0.02 | < 5 | 0.025 | 0.023 | < 0.05 | 0.1 | 0.024 | 0.028 | 0.024 | 0.025 | [7 empty] | 0.025375
| 1 | 5 | 0.03 | < 0.05 | 0.03 | < 5 | 0.021 | 0.028 | < 0.05 | 0.1 | 0.022 | 0.029 | 0.026 | 0.026 | [7 empty] | 90
| 1 | 6 | 0.03 | < 0.05 | 0.03 | < 5 | 0.024 | 0.024 | < 0.05 | 0.1 | 0.023 | 0.024 | 0.025 | 0.02
| Mean |  | 0.026667 |  | 0.03 |  | 0.023 | 0.024333 |  | 0.1 | 0.0235 | 0.028833 | 0.0245 | 0.024667
| Median |  | 0.03 |  | 0.03 |  | 0.023 | 0.024 |  | 0.1 | 0.023 | 0.0295 | 0.0245 | 0.0245
| Std Dev. |  | 0.005164 |  | 0.0063246 |  | 0.0014142 | 0.0020656 |  | 1.5202e-17 | 0.0013784 | 0.0027142 | 0.0010488 | 0.0029439
| Rel.Std.Dev. |  | 19.36% |  | 21.08% |  | 6.149% | 8.489% |  | 1.52e-14% | 5.866% | 9.413% | 4.281% | 11.93%
| PDM3 |  | 5.09% |  | 18.23% |  | -9.36% | -4.105% |  | 294.1% | -7.389% | 13.63% | -3.448% | -2.791%
| Z-Score (Absolute) |  | 0.75 | 0 | 2.25 | 1112.61 | 0.9 | 0.3 | 0 | NA | 0.67 | 1.72 | 0.22 | 0.15
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 250c (Certified Value 0.335 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.3397 | 0.35413 | 0.28966 | 0.42 | 0.336 | 0.34 | 0.3 | 0.34 | 0.35 | 0.34 | 0.32 | 0.34902 | 0.38 | 0.36 | 0.33 | 0.32 | 0.298 |  |  | 1
| 1 | 2 | 0.3342 | 0.34732 | 0.30315 | 0.43 | 0.345 | 0.32 | 0.3 | 0.35 | 0.34 | 0.33 | 0.33 | 0.36222 | 0.39 | 0.35 | 0.33 | 0.33 | 0.286 |  |  | #N/A
| 1 | 3 | 0.347 | 0.34704 | 0.31125 | 0.45 | 0.334 | 0.32 | 0.3 | 0.36 | 0.34 | 0.32 | 0.33 | 0.36955 | 0.38 | 0.36 | 0.32 | 0.33 | 0.289 |  |  | 16
| 1 | 4 | 0.3533 | 0.35426 | 0.30806 | 0.44 | 0.345 | 0.33 | 0.31 | 0.35 | 0.35 | 0.32 | 0.32 | 0.36213 | 0.4 | 0.36 | 0.33 | 0.33 | 0.289 |  |  | 0.33493
| 1 | 5 | 0.3465 | 0.34482 | 0.29225 | 0.44 | 0.339 | 0.34 | 0.31 | 0.35 | 0.35 | 0.34 | 0.32 | 0.36942 | 0.39 | 0.35 | 0.33 | 0.33 | 0.286 |  |  | 91
| 1 | 6 | 0.3402 | 0.34542 | 0.29553 | 0.44 | 0.335 | 0.34 | 0.3 | 0.35 | 0.35 | 0.3 | 0.32 | 0.34921 | 0.39 | 0.35 | 0.33 | 0.33 | 0.287
| Mean |  | 0.34348 | 0.34883 | 0.29998 | 0.43667 | 0.339 | 0.33167 | 0.30333 | 0.35 | 0.34667 | 0.325 | 0.32333 | 0.36026 | 0.38833 | 0.355 | 0.32833 | 0.32833 | 0.28917
| Median |  | 0.34335 | 0.34718 | 0.29934 | 0.44 | 0.3375 | 0.335 | 0.3 | 0.35 | 0.35 | 0.325 | 0.32 | 0.36217 | 0.39 | 0.355 | 0.33 | 0.33 | 0.288
| Std Dev. |  | 0.0067692 | 0.0042612 | 0.0088139 | 0.010328 | 0.0049396 | 0.0098319 | 0.005164 | 0.0063246 | 0.005164 | 0.015166 | 0.005164 | 0.0092329 | 0.0075277 | 0.0054772 | 0.0040825 | 0.0040825 | 0.004535
| Rel.Std.Dev. |  | 1.971% | 1.222% | 2.938% | 2.365% | 1.457% | 2.964% | 1.702% | 1.807% | 1.49% | 4.666% | 1.597% | 2.563% | 1.938% | 1.543% | 1.243% | 1.243% | 1.568%
| PDM3 |  | 2.552% | 4.149% | -10.44% | 30.37% | 1.214% | -0.9757% | -9.435% | 4.498% | 3.503% | -2.966% | -3.464% | 7.56% | 15.94% | 5.991% | -1.971% | -1.971% | -13.66%
| Z-Score (Absolute) |  | 0.22 | 0.47 | 1.88 | 4.7 | 0 | 0.35 | 1.72 | 0.53 | 0.37 | 0.67 | 0.75 | 1.02 | 2.38 | 0.77 | 0.51 | 0.51 | 2.4
| Analytical results for La in OREAS 250c (Certified Value 23.2 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 24.134 | 22.3638 | 23.5 | 24.1 | 23 | 21.8567 | 23.1 | 23 | 23.6 | 23.1643 | 17.8 | 23.7 | 25.2 | 22.4 | 21.6 |  |  |  |  | 1
| 1 | 2 | 23.456 | 22.014 | 23.6 | 24.6 | 22 | 22.0767 | 22.6 | 23 | 23.6 | 24.1932 | 18.3 | 23.7 | 27 | 23 | 21.6 |  |  |  |  | 23
| 1 | 3 | 23.671 | 22.6251 | 23.5 | 23.9 | 23 | 22.5433 | 22.5 | 22 | 23.8 | 24.1942 | 18 | 24.2 | 24.5 | 22.8 | 21.6 |  |  |  |  | 16
| 1 | 4 | 24.255 | 22.5537 | 23 | 22.9 | 23 | 22.2 | 23.2 | 22 | 22.9 | 23.9382 | 17.6 | 24.4 | 24.9 | 22.8 | 21.7 |  |  |  |  | 23.162
| 1 | 5 | 23.787 | 22.1324 | 23.4 | 24.6 | 23 | 22.0133 | 22.2 | 23 | 23.7 | 23.981 | 18.1 | 23.2 | 24.6 | 23.1 | 21.1 |  |  |  |  | 92
| 1 | 6 | 24.092 | 22.0894 | 23.2 | 25.5 | 23 | 22.0767 | 23.5 | 21 | 23.4 | 23.2966 | 17.7 | 23.7 | 24.8 | 23 | 21.6
| Mean |  | 23.8992 | 22.2964 | 23.3667 | 24.2667 | 22.8333 | 22.1278 | 22.85 | 22.3333 | 23.5 | 23.7946 | 17.9167 | 23.8167 | 25.1667 | 22.85 | 21.5333
| Median |  | 23.9395 | 22.2481 | 23.45 | 24.35 | 23 | 22.0767 | 22.85 | 22.5 | 23.6 | 23.9596 | 17.9 | 23.7 | 24.85 | 22.9 | 21.6
| Std Dev. |  | 0.30983 | 0.25625 | 0.22509 | 0.86872 | 0.40825 | 0.23226 | 0.49295 | 0.8165 | 0.32249 | 0.45147 | 0.26394 | 0.42622 | 0.93095 | 0.251 | 0.21602
| Rel.Std.Dev. |  | 1.296% | 1.149% | 0.9633% | 3.58% | 1.788% | 1.05% | 2.157% | 3.656% | 1.372% | 1.897% | 1.473% | 1.79% | 3.699% | 1.098% | 1.003%
| PDM3 |  | 3.183% | -3.737% | 0.8837% | 4.769% | -1.419% | -4.465% | -1.347% | -3.578% | 1.459% | 2.731% | -22.65% | 2.826% | 8.655% | -1.347% | -7.032%
| Z-Score (Absolute) |  | 0.98 | 0.52 | 0.48 | 1.32 | 0.02 | 0.67 | 0 | 0.48 | 0.61 | 0.88 | 4.61 | 0.9 | 2.16 | 0 | 1.23
| Analytical results for Li in OREAS 250c (Certified Value 7.61 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 7.6 | 8.1843 | 8 | 8.3 | 8.1 | 7.8 | 7.4012 | 6.9 | 6.6 | 3.8795 | 8.8 | 7.7 | 7.1 | 6.8 | 10 |  |  |  |  | 1
| 1 | 2 | 7.3 | 8.2233 | 8 | 8.6 | 8.2 | 7 | 7.5189 | 7.2 | 6.8 | 3.8716 | 8.7 | 7.5 | 6.9 | 7.2 | 9 |  |  |  |  | #N/A
| 1 | 3 | 7.4 | 8.2447 | 8 | 8.3 | 8.1 | 7.1 | 7.659 | 7 | 6.8 | 4.0511 | 8.5 | 7.6 | 6.9 | 7 | 11 |  |  |  |  | 16
| 1 | 4 | 7.8 | 8.4461 | 9 | 8.5 | 7.9 | 7.9 | 7.3591 | 7 | 6.7 | 4.4157 | 8.7 | 7.7 | 7.2 | 7 | 9 |  |  |  |  | 7.6061
| 1 | 5 | 7.5 | 8.0803 | 24 | 8.3 | 8.3 | 7.8 | 7.4882 | 7.2 | 6.8 | 4.5393 | 8.5 | 7.3 | 6.9 | 7.2 | 9 |  |  |  |  | 93
| 1 | 6 | 7.6 | 8.2782 | 11 | 8.4 | 8.5 | 7.4 | 7.4177 | 6.9 | 6.6 | 3.8343 | 8.7 | 7.5 | 6.9 | 7.1 | 9
| Mean |  | 7.5333 | 8.2428 | 11.3333 | 8.4 | 8.1833 | 7.5 | 7.474 | 7.0333 | 6.7167 | 4.0986 | 8.65 | 7.55 | 6.9833 | 7.05 | 9.5
| Median |  | 7.55 | 8.234 | 8.5 | 8.35 | 8.15 | 7.6 | 7.4529 | 7 | 6.75 | 3.9653 | 8.7 | 7.55 | 6.9 | 7.05 | 9
| Std Dev. |  | 0.17512 | 0.12066 | 6.314 | 0.12649 | 0.20412 | 0.38987 | 0.10781 | 0.13663 | 0.098319 | 0.30543 | 0.12247 | 0.15166 | 0.13292 | 0.15166 | 0.83666
| Rel.Std.Dev. |  | 2.325% | 1.464% | 55.71% | 1.506% | 2.494% | 5.198% | 1.442% | 1.943% | 1.464% | 7.452% | 1.416% | 2.009% | 1.903% | 2.151% | 8.807%
| PDM3 |  | -0.957% | 8.371% | 49% | 10.44% | 7.589% | -1.395% | -1.737% | -7.531% | -11.69% | -46.11% | 13.72% | -0.7379% | -8.188% | -7.312% | 24.9%
| Z-Score (Absolute) |  | 0.04 | 0.97 | NA | 1.17 | 0.89 | 0 | 0.03 | 0.61 | 1.02 | 4.44 | 1.5 | 0.07 | 0.67 | 0.59 | NA
Results from laboratories 1.06 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 250c (Indicative Value 0.09 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.22 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [15 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.25g | [15 empty] | 3
| 1 | 1 | 0.091 | 0.097636 | < 0.1 | 0.08 | [15 empty] | 1
| 1 | 2 | 0.092 | 0.093841 | < 0.1 | 0.08 | [15 empty] | 5
| 1 | 3 | 0.09 | 0.09828 | < 0.1 | 0.08 | [15 empty] | 16
| 1 | 4 | 0.094 | 0.098022 | < 0.1 | 0.09 | [15 empty] | 0.090077
| 1 | 5 | 0.093 | 0.097663 | < 0.1 | 0.08 | [15 empty] | 11
| 1 | 6 | 0.093 | 0.092952 | < 0.1 | 0.08
| Mean |  | 0.092167 | 0.096399 |  | 0.081667
| Median |  | 0.0925 | 0.09765 |  | 0.08
| Std Dev. |  | 0.001472 | 0.0023547 |  | 0.0040825
| Rel.Std.Dev. |  | 1.597% | 2.443% |  | 4.999%
| PDM3 |  | 2.319% | 7.018% |  | -9.337%
| Z-Score (Absolute) |  | 0.48 | 0.87 | 3.38 | 0.48
| Analytical results for Mg in OREAS 250c (Certified Value 0.989 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.99 | 1.2378 | 1.0036 | 0.89555 | 1.24 | 1.03 | 1.18 | 0.88 | 0.96 | 0.96 | 1.05 | 0.92 | 1.0251 | 1.05 | 1.03 | 0.96 | 0.93 | 0.94 |  | 1
| 1 | 2 | 0.97 | 1.2163 | 1.0043 | 0.89958 | 1.27 | 1.06 | 1.2 | 0.88 | 0.96 | 0.99 | 1.02 | 0.93 | 1.0868 | 1.07 | 1 | 0.95 | 0.94 | 0.931 |  | #N/A
| 1 | 3 | 1.01 | 1.2202 | 1.0093 | 0.91493 | 1.31 | 1.02 | 1.15 | 0.88 | 0.99 | 0.98 | 1 | 0.93 | 1.0597 | 1.04 | 1 | 0.95 | 0.94 | 0.937 |  | 16
| 1 | 4 | 1.03 | 1.1718 | 1.0221 | 0.90594 | 1.43 | 1.06 | 1.12 | 0.9 | 0.95 | 0.98 | 1 | 0.91 | 1.0647 | 1.07 | 1.03 | 0.97 | 0.94 | 0.936 |  | 0.98857
| 1 | 5 | 1.01 | 1.1985 | 0.99954 | 0.8933 | 1.31 | 1.04 | 1.18 | 0.89 | 0.94 | 1 | 1.03 | 0.93 | 1.0716 | 1.05 | 0.99 | 0.94 | 0.95 | 0.933 |  | 94
| 1 | 6 | 0.99 | 1.2118 | 1.0024 | 0.89378 | 1.31 | 1.03 | 1.24 | 0.87 | 0.95 | 0.98 | 0.95 | 0.92 | 1.0429 | 1.06 | 0.99 | 0.94 | 0.94 | 0.935
| Mean |  | 1 | 1.2094 | 1.0069 | 0.90051 | 1.3117 | 1.04 | 1.1783 | 0.88333 | 0.95833 | 0.98167 | 1.0083 | 0.92333 | 1.0585 | 1.0567 | 1.0067 | 0.95167 | 0.94 | 0.93533
| Median |  | 1 | 1.214 | 1.004 | 0.89756 | 1.31 | 1.035 | 1.18 | 0.88 | 0.955 | 0.98 | 1.01 | 0.925 | 1.0622 | 1.055 | 1 | 0.95 | 0.94 | 0.9355
| Std Dev. |  | 0.020976 | 0.022398 | 0.008112 | 0.0084865 | 0.064627 | 0.016733 | 0.041191 | 0.010328 | 0.017224 | 0.013292 | 0.034303 | 0.008165 | 0.021783 | 0.012111 | 0.018619 | 0.01169 | 0.0063246 | 0.0031411
| Rel.Std.Dev. |  | 2.098% | 1.852% | 0.8057% | 0.9424% | 4.927% | 1.609% | 3.496% | 1.169% | 1.797% | 1.354% | 3.402% | 0.8843% | 2.058% | 1.146% | 1.85% | 1.228% | 0.6728% | 0.3358%
| PDM3 |  | 1.156% | 22.34% | 1.851% | -8.908% | 32.68% | 5.202% | 19.2% | -10.65% | -3.059% | -0.6988% | 1.999% | -6.6% | 7.071% | 6.888% | 1.83% | -3.733% | -4.914% | -5.386%
| Z-Score (Absolute) |  | 0.04 | 2.56 | 0.04 | 1.28 | 3.83 | 0.46 | 2.18 | 1.49 | 0.56 | 0.27 | 0.06 | 0.99 | 0.69 | 0.66 | 0.04 | 0.64 | 0.79 | 0.85
| Analytical results for Mn in OREAS 250c (Certified Value 0.033 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0334 | 0.036343 | 0.034238 | 0.028229 | 0.03 | 0.035 | 0.035 | 0.0311 | 0.0346 | 0.0332 | 0.0343 | 0.031 | 0.035102 | 0.0353 | 0.0348 | 0.0319 | 0.0312 | 0.0409 |  | 1
| 1 | 2 | 0.0329 | 0.036199 | 0.033367 | 0.029029 | 0.03 | 0.0362 | 0.036 | 0.0301 | 0.0347 | 0.0343 | 0.0332 | 0.0315 | 0.035906 | 0.0364 | 0.0335 | 0.0318 | 0.0318 | 0.0406 |  | 25
| 1 | 3 | 0.0343 | 0.036841 | 0.03402 | 0.029362 | 0.03 | 0.0349 | 0.037 | 0.0308 | 0.0352 | 0.0339 | 0.0325 | 0.0312 | 0.035268 | 0.0349 | 0.0339 | 0.0314 | 0.0314 | 0.0406 |  | 16
| 1 | 4 | 0.0348 | 0.035516 | 0.035035 | 0.029275 | 0.03 | 0.0362 | 0.0324 | 0.0313 | 0.035 | 0.0342 | 0.0325 | 0.0305 | 0.036015 | 0.0357 | 0.0343 | 0.032 | 0.0314 | 0.0406 |  | 0.033367
| 1 | 5 | 0.0342 | 0.036135 | 0.034246 | 0.028458 | 0.04 | 0.0355 | 0.0349 | 0.0311 | 0.0349 | 0.0347 | 0.0337 | 0.0309 | 0.035368 | 0.0358 | 0.0332 | 0.0319 | 0.0318 | 0.0404 |  | 95
| 1 | 6 | 0.0335 | 0.035044 | 0.034676 | 0.028462 | 0.03 | 0.0352 | 0.0365 | 0.0305 | 0.0349 | 0.0336 | 0.0306 | 0.0306 | 0.035516 | 0.036 | 0.0337 | 0.0313 | 0.0319 | 0.0404
| Mean |  | 0.03385 | 0.036013 | 0.034264 | 0.028803 | 0.031667 | 0.0355 | 0.0353 | 0.030817 | 0.034883 | 0.033983 | 0.0328 | 0.03095 | 0.035529 | 0.035683 | 0.0339 | 0.031717 | 0.031583 | 0.040583
| Median |  | 0.03385 | 0.036167 | 0.034242 | 0.028746 | 0.03 | 0.03535 | 0.0355 | 0.03095 | 0.0349 | 0.03405 | 0.03285 | 0.03095 | 0.035442 | 0.03575 | 0.0338 | 0.03185 | 0.0316 | 0.0406
| Std Dev. |  | 0.00070071 | 0.00063726 | 0.00057089 | 0.00047961 | 0.0040825 | 0.00057966 | 0.001642 | 0.00044907 | 0.0002137 | 0.00053448 | 0.0012837 | 0.00037283 | 0.00036191 | 0.00052694 | 0.00057619 | 0.00029269 | 0.00028577 | 0.00018348
| Rel.Std.Dev. |  | 2.07% | 1.77% | 1.666% | 1.665% | 12.89% | 1.633% | 4.651% | 1.457% | 0.6126% | 1.573% | 3.914% | 1.205% | 1.019% | 1.477% | 1.7% | 0.9228% | 0.9048% | 0.4521%
| PDM3 |  | 1.447% | 7.929% | 2.687% | -13.68% | -5.096% | 6.392% | 5.793% | -7.644% | 4.544% | 1.847% | -1.7% | -7.244% | 6.48% | 6.941% | 1.597% | -4.946% | -5.346% | 21.63%
| Z-Score (Absolute) |  | 0.04 | 0.84 | 0.13 | 2.08 | 0.92 | 0.63 | 0.55 | 1.27 | 0.38 | 0.02 | 0.46 | 1.21 | 0.64 | 0.71 | 0.02 | 0.9 | 0.96 | 2.69
| Analytical results for Mo in OREAS 250c (Certified Value 1.32 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.4 | 1.3996 | < 10 | 1.5 | 1.4 | 1.08 | < 2 | 1.16 | 1 | 1.421 | 1.24 | 1.1877 | 1.36 | 1.24 | 1.47 | 1.21 | 1.37 |  |  | 1
| 1 | 2 | 1.4 | 1.3785 | < 10 | 1.5 | 1.39 | 1.06 | < 2 | 1.29 | 1.3 | 1.426 | 1.28 | 1.1695 | 1.4 | 1.28 | 1.51 | 1.18 | 1.43 |  |  | 26
| 1 | 3 | 1.4 | 1.3721 | < 10 | 1.5 | 1.36 | 1.06 | < 2 | 1.26 | 1.4 | 1.435 | 1.26 | 1.2746 | 1.32 | 1.27 | 1.45 | 1.22 | 1.36 |  |  | 16
| 1 | 4 | 1.4 | 1.3618 | < 10 | 1.5 | 1.29 | 1.1 | < 2 | 1.18 | 1.3 | 1.409 | 1.18 | 1.2575 | 1.33 | 1.26 | 1.45 | 1.22 | 1.38 |  |  | 1.3238
| 1 | 5 | 1.4 | 1.3434 | < 10 | 1.5 | 1.37 | 1.08 | < 2 | 1.22 | 1.3 | 1.462 | 1.22 | 1.1696 | 1.42 | 1.22 | 1.48 | 1.23 | 1.36 |  |  | 96
| 1 | 6 | 1.4 | 1.4064 | < 10 | 1.5 | 1.45 | 1.05 | < 2 | 1.2 | 1.3 | 1.463 | 1.22 | 1.273 | 1.4 | 1.24 | 1.45 | 1.22 | 1.44
| Mean |  | 1.4 | 1.377 |  | 1.5 | 1.3767 | 1.0717 |  | 1.2183 | 1.2667 | 1.436 | 1.2333 | 1.222 | 1.3717 | 1.2517 | 1.4683 | 1.2133 | 1.39
| Median |  | 1.4 | 1.3753 |  | 1.5 | 1.38 | 1.07 |  | 1.21 | 1.3 | 1.4305 | 1.23 | 1.2226 | 1.38 | 1.25 | 1.46 | 1.22 | 1.375
| Std Dev. |  | 2.4324e-16 | 0.023497 |  | 0 | 0.052789 | 0.018348 |  | 0.04916 | 0.13663 | 0.022181 | 0.035024 | 0.051612 | 0.041191 | 0.022286 | 0.024014 | 0.017512 | 0.035777
| Rel.Std.Dev. |  | 1.737e-14% | 1.706% |  | 0% | 3.835% | 1.712% |  | 4.035% | 10.79% | 1.545% | 2.84% | 4.224% | 3.003% | 1.781% | 1.635% | 1.443% | 2.574%
| PDM3 |  | 5.758% | 4.018% |  | 13.31% | 3.996% | -19.04% |  | -7.965% | -4.314% | 8.478% | -6.832% | -7.689% | 3.618% | -5.447% | 10.92% | -8.343% | 5.003%
| Z-Score (Absolute) |  | 0.16 | 0.03 | 20.39 | 0.72 | 0.03 | 1.69 | 2.09 | 0.86 | 0.59 | 0.36 | 0.78 | 0.84 | 0 | 0.67 | 0.54 | 0.89 | 0.1
| Analytical results for Na in OREAS 250c (Certified Value 0.351 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.36 | 0.37596 | 0.34194 | 1.19 | 0.36 | 0.383 | 0.32 | 0.33 | 0.35 | 0.347 | 0.33 | 0.31018 | 0.39 | 0.35 | 0.35 | 0.34 | 0.292 |  |  | 1
| 1 | 2 | 0.35 | 0.37316 | 0.36295 | 1.28 | 0.36 | 0.373 | 0.31 | 0.34 | 0.37 | 0.331 | 0.34 | 0.34961 | 0.4 | 0.34 | 0.35 | 0.35 | 0.277 |  |  | #N/A
| 1 | 3 | 0.37 | 0.36356 | 0.37325 | 1.32 | 0.36 | 0.369 | 0.31 | 0.34 | 0.36 | 0.319 | 0.34 | 0.33735 | 0.39 | 0.34 | 0.35 | 0.34 | 0.279 |  |  | 16
| 1 | 4 | 0.37 | 0.37041 | 0.36478 | 1.3 | 0.36 | 0.347 | 0.33 | 0.34 | 0.35 | 0.32 | 0.33 | 0.33889 | 0.4 | 0.35 | 0.36 | 0.34 | 0.285 |  |  | 0.35057
| 1 | 5 | 0.37 | 0.36498 | 0.34623 | 1.32 | 0.36 | 0.388 | 0.32 | 0.34 | 0.37 | 0.332 | 0.33 | 0.33906 | 0.39 | 0.34 | 0.35 | 0.35 | 0.279 |  |  | 97
| 1 | 6 | 0.36 | 0.37005 | 0.35056 | 1.3 | 0.36 | 0.386 | 0.31 | 0.33 | 0.36 | 0.293 | 0.33 | 0.31039 | 0.39 | 0.34 | 0.36 | 0.34 | 0.28
| Mean |  | 0.36333 | 0.36969 | 0.35662 | 1.285 | 0.36 | 0.37433 | 0.31667 | 0.33667 | 0.36 | 0.32367 | 0.33333 | 0.33091 | 0.39333 | 0.34333 | 0.35333 | 0.34333 | 0.282
| Median |  | 0.365 | 0.37023 | 0.35676 | 1.3 | 0.36 | 0.378 | 0.315 | 0.34 | 0.36 | 0.3255 | 0.33 | 0.33812 | 0.39 | 0.34 | 0.35 | 0.34 | 0.2795
| Std Dev. |  | 0.008165 | 0.0047264 | 0.012193 | 0.048888 | 6.0809e-17 | 0.015332 | 0.008165 | 0.005164 | 0.0089443 | 0.018129 | 0.005164 | 0.016567 | 0.005164 | 0.005164 | 0.005164 | 0.005164 | 0.0055857
| Rel.Std.Dev. |  | 2.247% | 1.278% | 3.419% | 3.804% | 1.689e-14% | 4.096% | 2.578% | 1.534% | 2.485% | 5.601% | 1.549% | 5.006% | 1.313% | 1.504% | 1.462% | 1.504% | 1.981%
| PDM3 |  | 3.641% | 5.453% | 1.725% | 266.5% | 2.69% | 6.779% | -9.671% | -3.966% | 2.69% | -7.674% | -4.917% | -5.608% | 12.2% | -2.064% | 0.7883% | -2.064% | -19.56%
| Z-Score (Absolute) |  | 0.4 | 0.66 | 0.13 | 37.69 | 0.27 | 0.85 | 1.48 | 0.67 | 0.27 | 1.2 | 0.81 | 0.91 | 1.62 | 0.4 | 0 | 0.4 | 2.89
| Analytical results for Nb in OREAS 250c (Certified Value 0.59 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.77 | 0.5879 | 13 | 1 | 0.3 | < 1 | < 5 | 0.69 | 0.55 | 1.3266 | 0.9 | 0.71 | 1.12 | 0.5 | 0.44 |  |  |  |  | 1
| 1 | 2 | 0.72 | 0.599 | 14 | 1.1 | 0.3 | < 1 | < 5 | 0.53 | 0.5 | 1.2077 | 0.9 | 0.67 | 1.02 | 0.5 | 0.45 |  |  |  |  | 27
| 1 | 3 | 0.66 | 0.59769 | 20 | 1.1 | 0.4 | < 1 | < 5 | 0.64 | 0.55 | 1.5178 | 1 | 0.71 | 0.93 | 0.49 | 0.5 |  |  |  |  | 16
| 1 | 4 | 0.74 | 0.57143 | 19 | 1.1 | 0.4 | < 1 | < 5 | 0.65 | 0.48 | 1.3455 | 0.8 | 0.63 | 1.04 | 0.51 | 0.47 |  |  |  |  | 0.58659
| 1 | 5 | 0.69 | 0.57914 | 26 | 1.2 | 0.4 | < 1 | < 5 | 0.53 | 0.52 | 1.2031 | 0.9 | 0.66 | 1.2 | 0.48 | 0.51 |  |  |  |  | 98
| 1 | 6 | 0.73 | 0.5815 | 24 | 1 | 0.4 | < 1 | < 5 | 0.7 | 0.54 | 1.5802 | 0.9 | 0.71 | 1.05 | 0.49 | 0.5
| Mean |  | 0.71833 | 0.58611 | 19.3333 | 1.0833 | 0.36667 |  |  | 0.62333 | 0.52333 | 1.3635 | 0.9 | 0.68167 | 1.06 | 0.495 | 0.47833
| Median |  | 0.725 | 0.5847 | 19.5 | 1.1 | 0.4 |  |  | 0.645 | 0.53 | 1.336 | 0.9 | 0.69 | 1.045 | 0.495 | 0.485
| Std Dev. |  | 0.038687 | 0.010852 | 5.2026 | 0.075277 | 0.05164 |  |  | 0.075807 | 0.028752 | 0.1565 | 0.063246 | 0.033714 | 0.091869 | 0.010488 | 0.029269
| Rel.Std.Dev. |  | 5.386% | 1.852% | 26.91% | 6.949% | 14.08% |  |  | 12.16% | 5.494% | 11.48% | 7.027% | 4.946% | 8.667% | 2.119% | 6.119%
| PDM3 |  | 22.46% | -0.08125% | 3196% | 84.68% | -37.49% |  |  | 6.264% | -10.78% | 132.4% | 53.43% | 16.21% | 80.71% | -15.61% | -18.45%
| Z-Score (Absolute) |  | 0.52 | 0.2 | NA | NA | NA | 0.67 | 10.26 | 0 | 0.55 | 4.05 | NA | 0.32 | 2.39 | 0.7 | 0.79
Results from laboratories 1.06, 1.07, 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 250c (Indicative Value 22.2 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.09 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 25g | 0.5g | [17 empty] | 1
| 1 | 1 | 22.604 | 21.9 | [17 empty] | 1
| 1 | 2 | 21.807 | 22.5 | [17 empty] | 6
| 1 | 3 | 22.434 | 21.4 | [17 empty] | 16
| 1 | 4 | 22.274 | 21.6 | [17 empty] | 22.1759
| 1 | 5 | 22.17 | 21.4 | [17 empty] | 12
| 1 | 6 | 22.822 | 23.2
| Mean |  | 22.3518 | 22
| Median |  | 22.354 | 21.75
| Std Dev. |  | 0.35422 | 0.71833
| Rel.Std.Dev. |  | 1.585% | 3.265%
| PDM3 |  | 0.7933% | -0.7933%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 250c (Certified Value 73 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 1
| 1 | 1 | 71.7 | 76.0428 | 71.9922 | 66.398 | 63 | 71 | 77.8 | 72.2 | 67.935 | 71.2 | 58 | 75.5 | 73.1744 | 75.2 | 74.5 | 73.6 | 72.2 | 58.7 |  | 1
| 1 | 2 | 71.2 | 78.0192 | 70.0603 | 65.932 | 65 | 71 | 76.8 | 71.5 | 68.4027 | 73.6 | 71 | 74.6 | 77.0296 | 76.6 | 73.8 | 78.3 | 72 | 59.2 |  | 29
| 1 | 3 | 73.9 | 77.1012 | 71.8334 | 68.764 | 67 | 71 | 78.4 | 72.4 | 69.306 | 72.7 | 71.4 | 76.3 | 77.2882 | 75.1 | 75.5 | 72.2 | 73.7 | 59.6 |  | 16
| 1 | 4 | 74.7 | 75.276 | 72.8991 | 68.225 | 65 | 71 | 77.6 | 74.1 | 68.9507 | 72.2 | 66.7 | 72.6 | 75.6749 | 74.7 | 75.1 | 75.7 | 73.1 | 59.5 |  | 72.6347
| 1 | 5 | 74 | 76.7664 | 72.2461 | 66.88 | 84 | 71 | 77.4 | 72.6 | 68.6607 | 71.4 | 67.9 | 76.1 | 76.0447 | 76 | 72.8 | 74.3 | 72.8 | 58.7 |  | 99
| 1 | 6 | 72 | 75.924 | 72.2896 | 67.007 | 80 | 71 | 77.3 | 71.2 | 68.5543 | 72.8 | 68.4 | 73.8 | 73.4782 | 76 | 72.8 | 73.4 | 73.1 | 59.1
| Mean |  | 72.9167 | 76.5216 | 71.8868 | 67.201 | 70.6667 | 71 | 77.55 | 72.3333 | 68.6349 | 72.3167 | 67.2333 | 74.8167 | 75.4483 | 75.6 | 74.0833 | 74.5833 | 72.8167 | 59.1333
| Median |  | 72.95 | 76.4046 | 72.1191 | 66.9435 | 66 | 71 | 77.5 | 72.3 | 68.6075 | 72.45 | 68.15 | 75.05 | 75.8598 | 75.6 | 74.15 | 73.95 | 72.95 | 59.15
| Std Dev. |  | 1.4552 | 0.9774 | 0.96603 | 1.0851 | 8.9592 | 0 | 0.53572 | 1.0191 | 0.46926 | 0.90866 | 4.8763 | 1.4359 | 1.7518 | 0.71274 | 1.1479 | 2.1535 | 0.63061 | 0.38297
| Rel.Std.Dev. |  | 1.996% | 1.277% | 1.344% | 1.615% | 12.68% | 0% | 0.6908% | 1.409% | 0.6837% | 1.257% | 7.253% | 1.919% | 2.322% | 0.9428% | 1.549% | 2.887% | 0.866% | 0.6476%
| PDM3 |  | 0.3883% | 5.351% | -1.03% | -7.481% | -2.709% | -2.251% | 6.767% | -0.4149% | -5.507% | -0.4378% | -7.436% | 3.004% | 3.874% | 4.083% | 1.994% | 2.683% | 0.2506% | -18.59%
| Z-Score (Absolute) |  | 0.11 | 1.25 | 0.22 | 1.71 | 0.61 | 0.5 | 1.58 | 0.08 | 1.25 | 0.08 | 1.7 | 0.71 | 0.91 | 0.96 | 0.48 | 0.64 | 0.08 | 4.27
| Analytical results for P in OREAS 250c (Certified Value 0.058 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0557 | 0.060125 | 0.053726 | 0.07 | 0.0624 | 0.06 | 0.0547 | 0.0528 | 0.057 | 0.059 | 0.056 | 0.062044 | 0.06 | 0.06 | 0.056 | 0.056 | 0.069 |  |  | 1
| 1 | 2 | 0.0547 | 0.059355 | 0.055882 | 0.07 | 0.0616 | 0.058 | 0.0545 | 0.0529 | 0.059 | 0.057 | 0.056 | 0.066254 | 0.06 | 0.058 | 0.057 | 0.056 | 0.068 |  |  | #N/A
| 1 | 3 | 0.0574 | 0.060264 | 0.056824 | 0.07 | 0.0621 | 0.058 | 0.0547 | 0.0534 | 0.06 | 0.057 | 0.057 | 0.065156 | 0.06 | 0.059 | 0.056 | 0.056 | 0.068 |  |  | 16
| 1 | 4 | 0.0576 | 0.06159 | 0.055435 | 0.07 | 0.0626 | 0.058 | 0.0552 | 0.0526 | 0.059 | 0.058 | 0.055 | 0.065135 | 0.06 | 0.059 | 0.057 | 0.056 | 0.071 |  |  | 0.05846
| 1 | 5 | 0.0563 | 0.060532 | 0.054219 | 0.07 | 0.0618 | 0.059 | 0.0544 | 0.0525 | 0.06 | 0.058 | 0.056 | 0.064081 | 0.06 | 0.058 | 0.058 | 0.056 | 0.068 |  |  | 100
| 1 | 6 | 0.0555 | 0.061099 | 0.052454 | 0.07 | 0.061 | 0.059 | 0.054 | 0.0526 | 0.059 | 0.054 | 0.056 | 0.062822 | 0.06 | 0.058 | 0.057 | 0.055 | 0.068
| Mean |  | 0.0562 | 0.060494 | 0.054757 | 0.07 | 0.061917 | 0.058667 | 0.054583 | 0.0528 | 0.059 | 0.057167 | 0.056 | 0.064249 | 0.06 | 0.058667 | 0.056833 | 0.055833 | 0.068667
| Median |  | 0.056 | 0.060398 | 0.054827 | 0.07 | 0.06195 | 0.0585 | 0.0546 | 0.0527 | 0.059 | 0.0575 | 0.056 | 0.064608 | 0.06 | 0.0585 | 0.057 | 0.056 | 0.068
| Std Dev. |  | 0.0011314 | 0.00078156 | 0.0015909 | 0 | 0.00058109 | 0.0008165 | 0.00039707 | 0.00032863 | 0.0010954 | 0.0017224 | 0.00063246 | 0.0015845 | 0 | 0.0008165 | 0.00075277 | 0.00040825 | 0.0012111
| Rel.Std.Dev. |  | 2.013% | 1.292% | 2.905% | 0% | 0.9385% | 1.392% | 0.7275% | 0.6224% | 1.857% | 3.013% | 1.129% | 2.466% | 0% | 1.392% | 1.325% | 0.7312% | 1.764%
| PDM3 |  | -3.867% | 3.479% | -6.336% | 19.74% | 5.912% | 0.3529% | -6.632% | -9.682% | 0.9231% | -2.213% | -4.209% | 9.901% | 2.634% | 0.3529% | -2.783% | -4.494% | 17.46%
| Z-Score (Absolute) |  | 0.62 | 0.46 | 0.99 | 2.87 | 0.82 | 0 | 1.03 | 1.48 | 0.08 | 0.38 | 0.67 | 1.41 | 0.34 | 0 | 0.46 | 0.72 | 2.53
| Analytical results for Pb in OREAS 250c (Certified Value 7.08 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 7.5 | 7.3969 | < 100 | 12 | 7.5 | 6.8 | 6.4 | 6.9957 | 6.6 | 5.8 | 4.431 | 6 | 9.6183 | 7.8 | 7.1 | 7.8 | 6.8 |  |  | 1
| 1 | 2 | 7.2 | 7.1835 | < 100 | 51 | 7.5 | 6.6 | 6.7 | 6.894 | 7 | 7 | 4.3932 | 6.1 | 8.5367 | 8.4 | 7.2 | 7.8 | 6.9 |  |  | 30
| 1 | 3 | 7.3 | 7.485 | < 100 | 20 | 7.6 | 7.6 | 7.2 | 7.0135 | 6.6 | 7.2 | 4.7947 | 6.1 | 9.721 | 7.9 | 7.5 | 7.7 | 6.9 |  |  | 16
| 1 | 4 | 7.5 | 7.2418 | < 100 | 15 | 7.3 | 7 | 6.5 | 7.1823 | 6.6 | 6.6 | 4.4206 | 6.1 | 8.9438 | 8 | 7.4 | 7.9 | 6.9 |  |  | 7.0819
| 1 | 5 | 7.4 | 7.3707 | < 100 | 24 | 7.4 | 7.1 | 6.5 | 7.0364 | 6.6 | 7 | 4.4028 | 6 | 7.7862 | 8.2 | 7.2 | 7.7 | 7 |  |  | 101
| 1 | 6 | 7.5 | 7.4804 | < 100 | 31 | 7.4 | 7 | 6.5 | 7.0488 | 6.8 | 6.6 | 4.2655 | 6 | 8.3248 | 8.1 | 7.1 | 7.4 | 6.7
| Mean |  | 7.4 | 7.3597 |  | 25.5 | 7.45 | 7.0167 | 6.6333 | 7.0285 | 6.7 | 6.7 | 4.4513 | 6.05 | 8.8218 | 8.0667 | 7.25 | 7.7167 | 6.8667
| Median |  | 7.45 | 7.3838 |  | 22 | 7.45 | 7 | 6.5 | 7.025 | 6.6 | 6.8 | 4.4117 | 6.05 | 8.7402 | 8.05 | 7.2 | 7.75 | 6.9
| Std Dev. |  | 0.12649 | 0.12388 |  | 14.181 | 0.10488 | 0.33714 | 0.29439 | 0.093274 | 0.16733 | 0.502 | 0.17861 | 0.054772 | 0.75613 | 0.21602 | 0.16432 | 0.17224 | 0.10328
| Rel.Std.Dev. |  | 1.709% | 1.683% |  | 55.61% | 1.408% | 4.805% | 4.438% | 1.327% | 2.497% | 7.492% | 4.012% | 0.9053% | 8.571% | 2.678% | 2.266% | 2.232% | 1.504%
| PDM3 |  | 4.492% | 3.923% |  | 260.1% | 5.198% | -0.9213% | -6.334% | -0.7547% | -5.393% | -5.393% | -37.15% | -14.57% | 24.57% | 13.91% | 2.374% | 8.963% | -3.039%
| Z-Score (Absolute) |  | 0.4 | 0.34 | 65.8 | NA | 0.48 | 0.19 | 0.78 | 0.17 | 0.67 | 0.67 | 4.13 | 1.67 | 2.58 | 1.42 | 0.17 | 0.89 | 0.42
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 250c (Indicative Value < 10 ppb) | [20 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.16 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [17 empty] | 1
|  |  | 25g | 01g | [17 empty] | 1
| 1 | 1 | < 10 | < 10 | [17 empty] | 1
| 1 | 2 | < 10 | < 10 | [17 empty] | 7
| 1 | 3 | < 10 | < 10 | [17 empty] | 16
| 1 | 4 | < 10 | < 10 | [17 empty] | < 10
| 1 | 5 | < 10 | < 10 | [17 empty] | 13
| 1 | 6 | < 10 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Pr in OREAS 250c (Indicative Value 5.77 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 5.939 | 5.5296 | 6 | [16 empty] | 1
| 1 | 2 | 5.919 | 5.2595 | 5.9 | [16 empty] | 8
| 1 | 3 | 5.863 | 5.4011 | 5.8 | [16 empty] | 16
| 1 | 4 | 5.9 | 5.6015 | 5.7 | [16 empty] | 5.7706
| 1 | 5 | 5.918 | 5.3591 | 6 | [16 empty] | 14
| 1 | 6 | 5.926 | 5.5549 | 6.3
| Mean |  | 5.9108 | 5.451 | 5.95
| Median |  | 5.9185 | 5.4654 | 5.95
| Std Dev. |  | 0.026634 | 0.13209 | 0.20736
| Rel.Std.Dev. |  | 0.4506% | 2.423% | 3.485%
| PDM3 |  | 2.43% | -5.539% | 3.109%
| Z-Score (Absolute) |  | 0 | 7.92 | 0.67
| Analytical results for Pt in OREAS 250c (Indicative Value < 5 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 1.01 | 1.16 | [17 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 25g | 01g | [17 empty] | 2
| 1 | 1 | < 5 | < 5 | [17 empty] | 1
| 1 | 2 | < 5 | < 5 | [17 empty] | 9
| 1 | 3 | < 5 | < 5 | [17 empty] | 16
| 1 | 4 | < 5 | < 5 | [17 empty] | < 5
| 1 | 5 | < 5 | < 5 | [17 empty] | 15
| 1 | 6 | < 5 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Rb in OREAS 250c (Certified Value 19.1 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 20.96 | 19.5231 | 17.2 | 18.4 | 18.1 | 17.6 | 20 | 17.9174 | 22.5 | 20.4 | 20.6 | 19 | 18.55 | [6 empty] | 1
| 1 | 2 | 19.98 | 18.8124 | 17.6 | 18.5 | 17.8 | 19.4 | 19.5 | 18.1923 | 23.3 | 20.2 | 20.7 | 18.7 | 18.44 | [6 empty] | 32
| 1 | 3 | 20.49 | 19.4845 | 16.9 | 18.1 | 18.2 | 19.2 | 20.5 | 19.0927 | 22.7 | 20.7 | 20.7 | 20 | 18.14 | [6 empty] | 16
| 1 | 4 | 21.05 | 19.6655 | 17.2 | 17.4 | 18.4 | 17.8 | 19.8 | 19.0268 | 22.9 | 20.3 | 20.9 | 19.5 | 18.5 | [6 empty] | 19.0763
| 1 | 5 | 20.73 | 19.1694 | 16.9 | 18.4 | 17.8 | 18.7 | 19.6 | 19.1321 | 22.9 | 19.6 | 20.7 | 19.6 | 17.97 | [6 empty] | 102
| 1 | 6 | 21.11 | 19.2905 | 16.8 | 18.8 | 18 | 18 | 19.6 | 17.5562 | 22.9 | 20 | 20.4 | 19.4 | 18.07
| Mean |  | 20.72 | 19.3242 | 17.1 | 18.2667 | 18.05 | 18.45 | 19.8333 | 18.4862 | 22.8667 | 20.2 | 20.6667 | 19.3667 | 18.2783
| Median |  | 20.845 | 19.3875 | 17.05 | 18.4 | 18.05 | 18.35 | 19.7 | 18.6095 | 22.9 | 20.25 | 20.7 | 19.45 | 18.29
| Std Dev. |  | 0.42877 | 0.30617 | 0.29665 | 0.48028 | 0.23452 | 0.75829 | 0.37238 | 0.68587 | 0.26583 | 0.37417 | 0.1633 | 0.45898 | 0.24766
| Rel.Std.Dev. |  | 2.069% | 1.584% | 1.735% | 2.629% | 1.299% | 4.11% | 1.878% | 3.71% | 1.163% | 1.852% | 0.7902% | 2.37% | 1.355%
| PDM3 |  | 8.617% | 1.3% | -10.36% | -4.244% | -5.38% | -3.283% | 3.969% | -3.093% | 19.87% | 5.891% | 8.337% | 1.522% | -4.183%
| Z-Score (Absolute) |  | 0.9 | 0 | 1.43 | 0.68 | 0.82 | 0.56 | 0.33 | 0.54 | 2.28 | 0.56 | 0.87 | 0.03 | 0.67
| Analytical results for Re in OREAS 250c (Certified Value < 0.001 ppm) | [20 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.001 | < 0.05 | 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | 0.002 | < 0.001 | [9 empty] | 1
| 1 | 2 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | 33
| 1 | 3 | < 0.001 | < 0.05 | 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | 0.003 | < 0.001 | [9 empty] | 16
| 1 | 4 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | [9 empty] | < 0.001
| 1 | 5 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | 0.001 | < 0.001 | < 0.05 | < 0.001 | 0.003 | < 0.001 | [9 empty] | 103
| 1 | 6 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.001
| Mean |  |  |  | 0.001 |  | 0.001 |  |  |  | 0.0026667
| Median |  |  |  | 0.001 |  | 0.001 |  |  |  | 0.003
| Std Dev. |  |  |  | 0 | [5 empty] | 0.00057735
| Rel.Std.Dev. |  |  |  | 0% | [5 empty] | 21.65%
PDM3
| Z-Score (Absolute) |  | 0.67 | 131.49 | 0.22 | 2.02 | 0.22 | 0.67 | 131.49 | 0.67 | 5.17 | 0.67
| Analytical results for S in OREAS 250c (Certified Value 0.024 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0244 | 0.02393 |  | 0.54 | 0.02 | < 0.001 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.023614 | 0.04 | 0.029242 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 |  | 1
| 1 | 2 | 0.0236 | 0.023338 |  | 0.57 | 0.02 | < 0.001 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.024313 | 0.04 | 0.026693 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02 |  | 18
| 1 | 3 | 0.0246 | 0.024164 |  | 0.58 | 0.02 | < 0.001 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.023613 | 0.04 | 0.02787 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02 |  | 16
| 1 | 4 | 0.025 | 0.024617 |  | 0.57 | 0.02 | < 0.001 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.023215 | 0.04 | 0.027994 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 |  | 0.02435
| 1 | 5 | 0.0252 | 0.024036 |  | 0.64 | 0.02 | < 0.001 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.023613 | 0.03 | 0.028555 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02 |  | 104
| 1 | 6 | 0.0242 | 0.02447 |  | 0.56 | 0.02 | < 0.001 | 0.03 | 0.02 | 0.02 | < 0.05 | 0.022816 | 0.04 | 0.026242 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02
| Mean |  | 0.0245 | 0.024092 |  | 0.57667 | 0.02 |  | 0.03 | 0.02 | 0.02 |  | 0.023531 | 0.038333 | 0.027766 | 0.03 | 0.02 | 0.026667 | 0.03 | 0.02
| Median |  | 0.0245 | 0.0241 |  | 0.57 | 0.02 |  | 0.03 | 0.02 | 0.02 |  | 0.023613 | 0.04 | 0.027932 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02
| Std Dev. |  | 0.00057619 | 0.00045194 |  | 0.033862 | 0 |  | 0 | 0 | 0 |  | 0.00049842 | 0.0040825 | 0.0011256 | 0 | 0 | 0.005164 | 0 | 0
| Rel.Std.Dev. |  | 2.352% | 1.876% |  | 5.872% | 0% |  | 0% | 0% | 0% |  | 2.118% | 10.65% | 4.054% | 0% | 0% | 19.36% | 0% | 0%
| PDM3 |  | 0.6142% | -1.06% |  | 2268% | -17.87% |  | 23.2% | -17.87% | -17.87% |  | -3.367% | 57.42% | 14.03% | 23.2% | -17.87% | 9.512% | 23.2% | -17.87%
| Z-Score (Absolute) |  | 0 | 0.06 | NA | 82.74 | 0.67 | 3.6 | 0.82 | 0.67 | 0.67 | 0.07 | 0.15 | 2.07 | 0.49 | 0.82 | 0.67 | 0.32 | 0.82 | 0.67
| Analytical results for Sb in OREAS 250c (Certified Value 0.45 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 0.69 | 0.4525 | < 10 | 1.7 | 0.46 | 0.47 | 0.35 | < 5 | 0.45 | 0.4 | 0.411 | 0.38 | 0.46384 | 0.58 | 0.44 | 0.54 | 0.44 | 0.49 |  | 1
| 1 | 2 | 0.61 | 0.43393 | < 10 | 1.6 | 0.49 | 0.45 | 0.3 | < 5 | 0.46 | 0.5 | 0.413 | 0.35 | 0.58486 | 0.62 | 0.44 | 0.5 | 0.46 | 0.48 |  | 34
| 1 | 3 | 0.63 | 0.43858 | < 10 | 4.1 | 0.42 | 0.47 | 0.31 | < 5 | 0.45 | 0.5 | 0.412 | 0.36 | 0.50393 | 0.72 | 0.46 | 0.52 | 0.46 | 0.47 |  | 16
| 1 | 4 | 0.64 | 0.45139 | < 10 | 2.6 | 0.49 | 0.53 | 0.35 | < 5 | 0.46 | 0.5 | 0.413 | 0.36 | 0.55868 | 0.62 | 0.44 | 0.52 | 0.47 | 0.48 |  | 0.44672
| 1 | 5 | 0.64 | 0.43271 | < 10 | 1.1 | 0.48 | 0.46 | 0.33 | < 5 | 0.44 | 0.5 | 0.412 | 0.38 | 0.47749 | 0.73 | 0.46 | 0.54 | 0.45 | 0.49 |  | 105
| 1 | 6 | 0.62 | 0.45431 | 20.891 | 2.5 | 0.45 | 0.45 | 0.3 | < 5 | 0.48 | 0.5 | 0.4135 | 0.38 | 0.54724 | 0.7 | 0.45 | 0.5 | 0.46 | 0.49
| Mean |  | 0.63833 | 0.4439 | 20.891 | 2.2667 | 0.465 | 0.47167 | 0.32333 |  | 0.45667 | 0.48333 | 0.41242 | 0.36833 | 0.52267 | 0.66167 | 0.44833 | 0.52 | 0.45667 | 0.48333
| Median |  | 0.635 | 0.44499 | 20.891 | 2.1 | 0.47 | 0.465 | 0.32 |  | 0.455 | 0.5 | 0.4125 | 0.37 | 0.52558 | 0.66 | 0.445 | 0.52 | 0.46 | 0.485
| Std Dev. |  | 0.027869 | 0.009915 |  | 1.0633 | 0.027386 | 0.029944 | 0.023381 |  | 0.013663 | 0.040825 | 0.00091742 | 0.013292 | 0.048211 | 0.062743 | 0.0098319 | 0.017889 | 0.010328 | 0.008165
| Rel.Std.Dev. |  | 4.366% | 2.234% |  | 46.91% | 5.889% | 6.349% | 7.231% |  | 2.992% | 8.447% | 0.2225% | 3.609% | 9.224% | 9.483% | 2.193% | 3.44% | 2.262% | 1.689%
| PDM3 |  | 42.89% | -0.6309% | 4577% | 407.4% | 4.092% | 5.584% | -27.62% |  | 2.226% | 8.196% | -7.679% | -17.55% | 17% | 48.12% | 0.3608% | 16.4% | 2.226% | 8.196%
| Z-Score (Absolute) |  | 2.16 | 0.31 | 91.35 | NA | 0.04 | 0.04 | 1.84 | 25.85 | 0.15 | NA | 0.71 | 1.27 | 0.69 | 2.46 | 0.25 | 0.66 | 0.15 | 0.19
Results from laboratories 1.06 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 250c (Certified Value 3.74 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 3.8 | 3.7837 | 4 | 3.9 | 3.5 | < 5 | 3.6 | 2.9 | 3.4 | 3.7868 | 4.3 | 4.1 | 3.9 | 3.8 | 3.4 |  |  |  |  | 1
| 1 | 2 | 3.7 | 3.7362 | 4 | 4 | 3.5 | < 5 | 3.7 | 3.5 | 3.4 | 3.7496 | 4.5 | 4.2 | 4.2 | 3.8 | 3.4 |  |  |  |  | 35
| 1 | 3 | 3.9 | 3.794 | 4 | 4 | 3.5 | < 5 | 3.6 | 3.6 | 3.5 | 3.438 | 4.4 | 4.2 | 3.8 | 3.9 | 3.4 |  |  |  |  | 16
| 1 | 4 | 4 | 3.8361 | 4 | 3.5 | 3.5 | < 5 | 3.6 | 3.3 | 3.4 | 3.4872 | 4.4 | 4.2 | 3.8 | 3.8 | 3.4 |  |  |  |  | 3.7374
| 1 | 5 | 3.9 | 3.7922 | 4 | 3.9 | 3.5 | < 5 | 3.6 | 3.4 | 3.5 | 3.537 | 4.3 | 4 | 3.7 | 3.8 | 3.4 |  |  |  |  | 106
| 1 | 6 | 3.8 | 3.7946 | 4 | 4.1 | 3.4 | < 5 | 3.7 | 3.3 | 3.4 | 3.2991 | 4.4 | 4 | 3.8 | 3.9 | 3.4
| Mean |  | 3.85 | 3.7895 | 4 | 3.9 | 3.4833 |  | 3.6333 | 3.3333 | 3.4333 | 3.5496 | 4.3833 | 4.1167 | 3.8667 | 3.8333 | 3.4
| Median |  | 3.85 | 3.7931 | 4 | 3.95 | 3.5 |  | 3.6 | 3.35 | 3.4 | 3.5121 | 4.4 | 4.15 | 3.8 | 3.8 | 3.4
| Std Dev. |  | 0.10488 | 0.031932 | 0 | 0.20976 | 0.040825 |  | 0.05164 | 0.24221 | 0.05164 | 0.18735 | 0.075277 | 0.098319 | 0.17512 | 0.05164 | 0
| Rel.Std.Dev. |  | 2.724% | 0.8426% | 0% | 5.379% | 1.172% |  | 1.421% | 7.266% | 1.504% | 5.278% | 1.717% | 2.388% | 4.529% | 1.347% | 0%
| PDM3 |  | 3.014% | 1.394% | 7.027% | 4.352% | -6.797% |  | -2.784% | -10.81% | -8.135% | -5.024% | 17.28% | 10.15% | 3.46% | 2.568% | -9.027%
| Z-Score (Absolute) |  | 0.45 | 0.25 | NA | 0.61 | 0.74 | 3.92 | 0.25 | 1.22 | 0.9 | 0.52 | 2.17 | 1.31 | 0.5 | 0.39 | 1.01
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 250c (Indicative Value 0.19 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | < 1.5 | < 1 | 0.3 | 0.3 | < 5 | 0.2 | < 0.5 | 0.058 | 0.2 | < 1 | 1.2 | < 0.2 | 0.2 | < 0.2 |  |  |  |  | 1
| 1 | 2 | < 1 | < 1.5 | < 1 | 0.3 | 0.3 | < 5 | < 0.2 | < 0.5 | 0.056 | 0.2 | < 1 | 1.3 | < 0.2 | < 0.2 | < 0.2 |  |  |  |  | 1
| 1 | 3 | < 1 | < 1.5 | < 1 | 0.3 | 0.3 | < 5 | 0.3 | < 0.5 | 0.058 | < 0.2 | < 1 | 0.9 | < 0.2 | < 0.2 | < 0.2 |  |  |  |  | 16
| 1 | 4 | < 1 | < 1.5 | < 1 | 0.2 | 0.3 | < 5 | 0.3 | < 0.5 | 0.059 | 0.2 | < 1 | 1.2 | < 0.2 | < 0.2 | < 0.2 |  |  |  |  | 0.19292
| 1 | 5 | < 1 | < 1.5 | < 1 | 0.2 | 0.3 | < 5 | < 0.2 | < 0.5 | 0.055 | 0.3 | < 1 | 1 | < 0.2 | < 0.2 | < 0.2 |  |  |  |  | 10
| 1 | 6 | < 1 | < 1.5 | < 1 | 0.4 | 0.2 | < 5 | 0.2 | < 0.5 | 0.059 | 0.3 | < 1 | 1 | < 0.2 | < 0.2 | < 0.2
| Mean |  |  |  |  | 0.28333 | 0.28333 |  | 0.25 |  | 0.0575 | 0.24 |  | 1.1 |  | 0.2
| Median |  |  |  |  | 0.3 | 0.3 |  | 0.25 |  | 0.058 | 0.2 |  | 1.1 |  | 0.2
| Std Dev. |  |  |  |  | 0.075277 | 0.040825 |  | 0.057735 |  | 0.0016432 | 0.054772 |  | 0.15492
| Rel.Std.Dev. |  |  |  |  | 26.57% | 14.41% |  | 23.09% |  | 2.858% | 22.82% |  | 14.08%
| PDM3 |  |  |  |  | 46.87% | 46.87% |  | 29.59% |  | -70.19% | 24.41% |  | 470.2% |  | 3.672%
| Z-Score (Absolute) |  | 0.8 | 1.72 | 0.8 | 0 | 0 | 8.15 | 0.31 | 0.12 | 0.83 | 0.25 | 0.8 | 3 | 0.67 | 0.61 | 0.67
| Analytical results for Sm in OREAS 250c (Indicative Value 4 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 4.154 | 4.0854 | 4.5 | [16 empty] | 1
| 1 | 2 | 4.005 | 3.8768 | 3.4 | [16 empty] | 5
| 1 | 3 | 4.059 | 3.9254 | 4.3 | [16 empty] | 16
| 1 | 4 | 4.035 | 4.0474 | 3 | [16 empty] | 3.9986
| 1 | 5 | 4.064 | 3.9871 | 4.8 | [16 empty] | 11
| 1 | 6 | 4.109 | 4.0262 | 3.6
| Mean |  | 4.071 | 3.9914 | 3.9333
| Median |  | 4.0615 | 4.0067 | 3.95
| Std Dev. |  | 0.05325 | 0.078349 | 0.70333
| Rel.Std.Dev. |  | 1.308% | 1.963% | 17.88%
| PDM3 |  | 1.811% | -0.1798% | -1.632%
| Z-Score (Absolute) |  | 0.92 | 0 | 0.67
| Analytical results for Sn in OREAS 250c (Certified Value 1.24 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 1.33 | < 2.5 | 1 | 1.6 | 1.45 | 1.2 | < 10 | 1.2 | 1.1 | 1.1924 | 1.6 | 1.2 | 1.4 | 1.2 | 1.1 |  |  |  |  | 1
| 1 | 2 | 1.22 | < 2.5 | 5 | 1.6 | 1.49 | 1.2 | < 10 | 1.2 | 1.1 | 1.3339 | 1.6 | 1.2 | 1.5 | 1.2 | 1 |  |  |  |  | 21
| 1 | 3 | 1.31 | < 2.5 | 2 | 1.5 | 1.48 | 1.2 | < 10 | 1.2 | 1.1 | 1.3008 | 1.6 | 1.3 | 1.4 | 1.2 | 1 |  |  |  |  | 16
| 1 | 4 | 1.32 | < 2.5 | 4 | 1.6 | 1.39 | 1.2 | < 10 | 1.2 | 1.1 | 1.223 | 1.6 | 1.2 | 1.4 | 1.2 | 1 |  |  |  |  | 1.2367
| 1 | 5 | 1.29 | < 2.5 | 7 | 1.6 | 1.44 | 1.2 | < 10 | 1.2 | 1.2 | 1.3339 | 1.6 | 1.2 | 1.4 | 1.2 | 1 |  |  |  |  | 107
| 1 | 6 | 1.28 | < 2.5 | 4 | 1.6 | 1.51 | 1.1 | < 10 | 1.2 | 1.1 | 1.2056 | 1.6 | 1.2 | 1.4 | 1.2 | 1
| Mean |  | 1.2917 |  | 3.8333 | 1.5833 | 1.46 | 1.1833 |  | 1.2 | 1.1167 | 1.2649 | 1.6 | 1.2167 | 1.4167 | 1.2 | 1.0167
| Median |  | 1.3 |  | 4 | 1.6 | 1.465 | 1.2 |  | 1.2 | 1.1 | 1.2619 | 1.6 | 1.2 | 1.4 | 1.2 | 1
| Std Dev. |  | 0.039707 |  | 2.137 | 0.040825 | 0.042895 | 0.040825 |  | 0 | 0.040825 | 0.065341 | 2.4324e-16 | 0.040825 | 0.040825 | 0 | 0.040825
| Rel.Std.Dev. |  | 3.074% |  | 55.75% | 2.578% | 2.938% | 3.45% |  | 0% | 3.656% | 5.166% | 1.52e-14% | 3.355% | 2.882% | 0% | 4.016%
| PDM3 |  | 4.448% |  | 210% | 28.03% | 18.06% | -4.312% |  | -2.964% | -9.703% | 2.286% | 29.38% | -1.617% | 14.56% | -2.964% | -17.79%
| Z-Score (Absolute) |  | 0.21 | 0.05 | NA | 2.04 | 1.27 | 0.47 | 23.48 | 0.36 | 0.88 | 0.05 | 2.15 | 0.26 | 1 | 0.36 | 1.51
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 250c (Certified Value 63 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  | 0
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 66.56 | 65.383 | 224 | 62 | 78.4 | 61.6 | 62.5133 | 64.1 | 63.7 | 65.2 | 57.3225 | 61.8 | 64.8 | 68 | 62.8 | 51.3 |  |  |  | 1
| 1 | 2 | 65.78 | 64.8502 | 231 | 63 | 78.5 | 60.7 | 63.4867 | 62.8 | 61.8 | 65.5 | 62.1929 | 62 | 63.7 | 70.8 | 62.8 | 49.8 |  |  |  | 22
| 1 | 3 | 66.08 | 65.1319 | 235 | 61.2 | 73.2 | 61.3 | 64.53 | 63.4 | 59.6 | 65.1 | 60.261 | 59.5 | 64.6 | 66.8 | 64.6 | 50.7 |  |  |  | 16
| 1 | 4 | 66.85 | 66.1046 | 232 | 62.4 | 72.4 | 62.6 | 63.5967 | 64.6 | 59.8 | 63.6 | 59.8201 | 62.5 | 65.1 | 69 | 63.3 | 50.7 |  |  |  | 63.4029
| 1 | 5 | 65.45 | 64.6996 | 238 | 60.9 | 77.3 | 61.7 | 63.4067 | 61.6 | 62.6 | 64.8 | 60.2406 | 61.2 | 62.7 | 66.3 | 63.9 | 50.1 |  |  |  | 108
| 1 | 6 | 66.88 | 65.6251 | 234 | 61.5 | 75.5 | 61.1 | 63.1533 | 64.9 | 56 | 64.5 | 56.5514 | 61.8 | 63.3 | 67.5 | 64.2 | 50.2
| Mean |  | 66.2667 | 65.2991 | 232.3333 | 61.8333 | 75.8833 | 61.5 | 63.4478 | 63.5667 | 60.5833 | 64.7833 | 59.3981 | 61.4667 | 64.0333 | 68.0667 | 63.6 | 50.4667
| Median |  | 66.32 | 65.2574 | 233 | 61.75 | 76.4 | 61.45 | 63.4467 | 63.75 | 60.8 | 64.95 | 60.0303 | 61.8 | 64.15 | 67.75 | 63.6 | 50.45
| Std Dev. |  | 0.59011 | 0.51982 | 4.761 | 0.78655 | 2.6332 | 0.64807 | 0.65673 | 1.234 | 2.7513 | 0.67355 | 2.0907 | 1.0501 | 0.94587 | 1.6367 | 0.751 | 0.53914
| Rel.Std.Dev. |  | 0.8905% | 0.7961% | 2.049% | 1.272% | 3.47% | 1.054% | 1.035% | 1.941% | 4.541% | 1.04% | 3.52% | 1.708% | 1.477% | 2.405% | 1.181% | 1.068%
| PDM3 |  | 4.517% | 2.991% | 266.4% | -2.476% | 19.68% | -3.001% | 0.07071% | 0.2582% | -4.447% | 2.177% | -6.317% | -3.054% | 0.9943% | 7.356% | 0.3108% | -20.4%
| Z-Score (Absolute) |  | 0.86 | 0.55 | 54.19 | 0.56 | 3.95 | 0.67 | 0.04 | 0.01 | 0.96 | 0.39 | 1.34 | 0.68 | 0.14 | 1.44 | 0.01 | 4.21
| Analytical results for Ta in OREAS 250c (Certified Value < 0.01 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | 0.03 | [6 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | 0.04 | [6 empty] | 23
| 1 | 3 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | 0.03 | [6 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | 0.03 | [6 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | 0.03 | [6 empty] | 109
| 1 | 6 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | 0.04
| Mean | [9 empty] | 0.3 |  | 0.01 |  | 0.033333
| Median | [9 empty] | 0.3 |  | 0.01 |  | 0.03
| Std Dev. | [9 empty] | 0 |  | 0 |  | 0.005164
| Rel.Std.Dev. | [9 empty] | 0% |  | 0% |  | 15.49%
PDM3
| Z-Score (Absolute) |  | 0.67 | 2.02 | 2.02 | 0.67 | 335.81 | 0.67 | 0.67 | 2.02 | 39.11 | 0.67 | 0 | 0.67 | 3.15
| Analytical results for Tb in OREAS 250c (Certified Value 0.45 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.17 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.25g | [14 empty] | 3
| 1 | 1 | 0.454 | 0.44363 | 0.4 | 0.5 | 0.43 | [14 empty] | 1
| 1 | 2 | 0.426 | 0.41997 | 0.4 | 0.5 | 0.44 | [14 empty] | 24
| 1 | 3 | 0.437 | 0.45319 | 0.4 | 0.5 | 0.43 | [14 empty] | 16
| 1 | 4 | 0.452 | 0.45709 | 0.4 | 0.5 | 0.44 | [14 empty] | 0.44776
| 1 | 5 | 0.448 | 0.44099 | 0.4 | 0.5 | 0.44 | [14 empty] | 110
| 1 | 6 | 0.457 | 0.444 | 0.5 | 0.5 | 0.42
| Mean |  | 0.44567 | 0.44315 | 0.41667 | 0.5 | 0.43333
| Median |  | 0.45 | 0.44381 | 0.4 | 0.5 | 0.435
| Std Dev. |  | 0.011877 | 0.012946 | 0.040825 | 0 | 0.008165
| Rel.Std.Dev. |  | 2.665% | 2.921% | 9.798% | 0% | 1.884%
| PDM3 |  | -0.4681% | -1.031% | -6.945% | 11.67% | -3.222%
| Z-Score (Absolute) |  | 0.17 | 0 | 1.82 | 3.91 | 0.67
| Analytical results for Te in OREAS 250c (Indicative Value 0.023 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 3
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.1 | < 0.1 | < 0.05 | 0.03 | 0.03 | < 5 | 0.02 | < 0.2 | < 0.03 | < 0.01 | < 0.05 | 0.03 | 0.02 | 0.05 | 0.02 |  |  |  |  | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.05 | 0.06 | 0.03 | < 5 | 0.01 | < 0.2 | < 0.03 | 0.01 | < 0.05 | 0.03 | 0.01 | 0.02 | 0.03 |  |  |  |  | 2
| 1 | 3 | < 0.1 | < 0.1 | < 0.05 | 0.03 | 0.02 | < 5 | 0.01 | < 0.2 | < 0.03 | < 0.01 | < 0.05 | 0.04 | 0.02 | 0.01 | 0.01 |  |  |  |  | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.05 | < 0.02 | 0.03 | < 5 | 0.02 | < 0.2 | < 0.03 | 0.01 | < 0.05 | 0.04 | 0.01 | 0.04 | 0.02 |  |  |  |  | 0.022604
| 1 | 5 | < 0.1 | < 0.1 | < 0.05 | 0.03 | 0.02 | < 5 | 0.02 | < 0.2 | < 0.03 | 0.01 | < 0.05 | 0.06 | 0.01 | 0.04 | 0.01 |  |  |  |  | 12
| 1 | 6 | < 0.1 | < 0.1 | < 0.05 | < 0.02 | 0.03 | < 5 | 0.02 | < 0.2 | < 0.03 | 0.01 | < 0.05 | 0.04 | 0.02 | < 0.01 | 0.02
| Mean |  |  |  |  | 0.0375 | 0.026667 |  | 0.016667 |  |  | 0.01 |  | 0.04 | 0.015 | 0.032 | 0.018333
| Median |  |  |  |  | 0.03 | 0.03 |  | 0.02 |  |  | 0.01 |  | 0.04 | 0.015 | 0.04 | 0.02
| Std Dev. |  |  |  |  | 0.015 | 0.005164 |  | 0.005164 |  |  | 0 |  | 0.010954 | 0.0054772 | 0.016432 | 0.0075277
| Rel.Std.Dev. |  |  |  |  | 40% | 19.36% |  | 30.98% |  |  | 0% |  | 27.39% | 36.51% | 51.35% | 41.06%
| PDM3 |  |  |  |  | 65.9% | 17.97% |  | -26.27% |  |  | -55.76% |  | 76.96% | -33.64% | 41.57% | -18.89%
| Z-Score (Absolute) |  | 1.35 | 1.35 | 0.1 | 0.1 | 0 | 142.95 | 0.58 | 4.24 | 0.67 | 1.06 | 0.1 | 0.77 | 0.67 | 0.05 | 0.48
| Analytical results for Th in OREAS 250c (Certified Value 5.85 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 6.14 | 5.7747 | < 100 | 5.86 | 6.1 | < 10 | 5.5 | 5.2 | 4.9 | 6.4109 | 5.69 | 5.7 | 6.4 | 5.5 | 6.08 |  |  |  |  | 1
| 1 | 2 | 5.88 | 5.6852 | < 100 | 6.04 | 6.3 | < 10 | 6.1 | 6 | 4.9 | 6.8773 | 5.95 | 5.7 | 6.8 | 5.6 | 6.13 |  |  |  |  | 26
| 1 | 3 | 5.96 | 5.8346 | < 100 | 5.62 | 6.4 | < 10 | 6 | 6.2 | 5 | 6.9523 | 5.62 | 5.8 | 6.3 | 5.7 | 6.09 |  |  |  |  | 16
| 1 | 4 | 6.09 | 5.6619 | < 100 | 5.92 | 5.7 | < 10 | 5.5 | 5.7 | 4.9 | 6.6836 | 5.71 | 5.8 | 6.4 | 5.8 | 6.09 |  |  |  |  | 5.8503
| 1 | 5 | 6.05 | 5.6061 | < 100 | 5.86 | 6.1 | < 10 | 5.9 | 5.9 | 4.8 | 6.772 | 5.46 | 5.7 | 6.2 | 5.7 | 6.05 |  |  |  |  | 111
| 1 | 6 | 6.18 | 5.9275 | < 100 | 5.68 | 6.4 | < 10 | 5.6 | 5.7 | 4.9 | 6.4914 | 5.69 | 5.8 | 6.1 | 5.5 | 6.09
| Mean |  | 6.05 | 5.7483 |  | 5.83 | 6.1667 |  | 5.7667 | 5.7833 | 4.9 | 6.6979 | 5.6867 | 5.75 | 6.3667 | 5.6333 | 6.0883
| Median |  | 6.07 | 5.73 |  | 5.86 | 6.2 |  | 5.75 | 5.8 | 4.9 | 6.7278 | 5.69 | 5.75 | 6.35 | 5.65 | 6.09
| Std Dev. |  | 0.11278 | 0.11994 |  | 0.15531 | 0.26583 |  | 0.26583 | 0.34303 | 0.063246 | 0.21335 | 0.15858 | 0.054772 | 0.24221 | 0.12111 | 0.025626
| Rel.Std.Dev. |  | 1.864% | 2.087% |  | 2.664% | 4.311% |  | 4.61% | 5.931% | 1.291% | 3.185% | 2.789% | 0.9526% | 3.804% | 2.15% | 0.4209%
| PDM3 |  | 3.413% | -1.743% |  | -0.3476% | 5.407% |  | -1.43% | -1.145% | -16.24% | 14.49% | -2.798% | -1.715% | 8.826% | -3.709% | 4.068%
| Z-Score (Absolute) |  | 0.67 | 0.09 | 111.81 | 0.12 | 0.97 | 1.98 | 0.04 | 0 | 2.23 | 2.31 | 0.24 | 0.08 | 1.48 | 0.38 | 0.77
| Analytical results for Ti in OREAS 250c (Certified Value 0.162 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.1637 | 0.17011 | 0.11824 | 0.17 | 0.205 | 0.22 | 0.16 | 0.15 | 0.154 | 0.166 | 0.154 | 0.25528 | 0.21 | 0.18 | 0.162 | 0.157 | 0.14 |  |  | 1
| 1 | 2 | 0.1602 | 0.16818 | 0.13264 | 0.18 | 0.203 | 0.21 | 0.16 | 0.16 | 0.161 | 0.161 | 0.157 | 0.26635 | 0.21 | 0.172 | 0.162 | 0.158 | 0.135 |  |  | 27
| 1 | 3 | 0.168 | 0.16912 | 0.13406 | 0.19 | 0.204 | 0.21 | 0.16 | 0.16 | 0.165 | 0.157 | 0.155 | 0.26002 | 0.2 | 0.176 | 0.16 | 0.161 | 0.138 |  |  | 16
| 1 | 4 | 0.1725 | 0.17296 | 0.13387 | 0.19 | 0.204 | 0.21 | 0.17 | 0.16 | 0.16 | 0.157 | 0.15 | 0.24778 | 0.21 | 0.176 | 0.165 | 0.157 | 0.137 |  |  | 0.1622
| 1 | 5 | 0.1661 | 0.16748 | 0.12386 | 0.2 | 0.201 | 0.21 | 0.17 | 0.15 | 0.164 | 0.165 | 0.155 | 0.26964 | 0.2 | 0.168 | 0.161 | 0.159 | 0.137 |  |  | 112
| 1 | 6 | 0.1647 | 0.1694 | 0.12519 | 0.2 | 0.203 | 0.21 | 0.16 | 0.15 | 0.163 | 0.145 | 0.153 | 0.24513 | 0.21 | 0.172 | 0.159 | 0.155 | 0.137
| Mean |  | 0.16587 | 0.16954 | 0.12798 | 0.18833 | 0.20333 | 0.21167 | 0.16333 | 0.155 | 0.16117 | 0.1585 | 0.154 | 0.25737 | 0.20667 | 0.174 | 0.1615 | 0.15783 | 0.13733
| Median |  | 0.1654 | 0.16926 | 0.12891 | 0.19 | 0.2035 | 0.21 | 0.16 | 0.155 | 0.162 | 0.159 | 0.1545 | 0.25765 | 0.21 | 0.174 | 0.1615 | 0.1575 | 0.137
| Std Dev. |  | 0.0041659 | 0.0019118 | 0.0065263 | 0.01169 | 0.0013663 | 0.0040825 | 0.005164 | 0.0054772 | 0.0039707 | 0.0076354 | 0.0023664 | 0.0098427 | 0.005164 | 0.0041952 | 0.0020736 | 0.0020412 | 0.001633
| Rel.Std.Dev. |  | 2.512% | 1.128% | 5.1% | 6.207% | 0.6719% | 1.929% | 3.162% | 3.534% | 2.464% | 4.817% | 1.537% | 3.824% | 2.499% | 2.411% | 1.284% | 1.293% | 1.189%
| PDM3 |  | 2.26% | 4.525% | -21.1% | 16.11% | 25.36% | 30.5% | 0.6984% | -4.439% | -0.6374% | -2.281% | -5.056% | 58.67% | 27.41% | 7.275% | -0.4319% | -2.693% | -15.33%
| Z-Score (Absolute) |  | 0.18 | 0.45 | 2.55 | 1.81 | 2.89 | 3.49 | 0 | 0.6 | 0.16 | 0.35 | 0.67 | 6.79 | 3.13 | 0.77 | 0.13 | 0.4 | 1.88
| Analytical results for Tl in OREAS 250c (Certified Value 0.14 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.15 | 0.13632 | 0.13 | 0.15 | 0.16 | < 10 | 0.12 | < 0.1 | 0.132 | 0.12 | 0.12031 | 0.17 | 0.12 | 0.14 | 0.12 |  |  |  |  | 1
| 1 | 2 | 0.13 | 0.12706 | 0.14 | 0.15 | 0.15 | < 10 | 0.13 | < 0.1 | 0.134 | 0.11 | 0.14464 | 0.17 | 0.12 | 0.15 | 0.13 |  |  |  |  | 28
| 1 | 3 | 0.14 | 0.13442 | 0.14 | 0.15 | 0.14 | < 10 | 0.12 | 0.1 | 0.135 | 0.12 | 0.11805 | 0.17 | 0.12 | 0.14 | 0.12 |  |  |  |  | 16
| 1 | 4 | 0.14 | 0.13591 | 0.14 | 0.14 | 0.17 | < 10 | 0.12 | 0.1 | 0.132 | 0.11 | 0.14038 | 0.16 | 0.12 | 0.15 | 0.12 |  |  |  |  | 0.1356
| 1 | 5 | 0.14 | 0.13326 | 0.14 | 0.15 | 0.16 | < 10 | 0.12 | < 0.1 | 0.135 | 0.11 | 0.13487 | 0.16 | 0.12 | 0.14 | 0.12 |  |  |  |  | 113
| 1 | 6 | 0.14 | 0.12936 | 0.13 | 0.16 | 0.15 | < 10 | 0.13 | < 0.1 | 0.131 | 0.11 | 0.1135 | 0.17 | 0.12 | 0.13 | 0.12
| Mean |  | 0.14 | 0.13272 | 0.13667 | 0.15 | 0.155 |  | 0.12333 | 0.1 | 0.13317 | 0.11333 | 0.12863 | 0.16667 | 0.12 | 0.14167 | 0.12167
| Median |  | 0.14 | 0.13384 | 0.14 | 0.15 | 0.155 |  | 0.12 | 0.1 | 0.133 | 0.11 | 0.12759 | 0.17 | 0.12 | 0.14 | 0.12
| Std Dev. |  | 0.0063246 | 0.0037314 | 0.005164 | 0.0063246 | 0.010488 |  | 0.005164 | 0 | 0.0017224 | 0.005164 | 0.012986 | 0.005164 | 0 | 0.0075277 | 0.0040825
| Rel.Std.Dev. |  | 4.518% | 2.811% | 3.779% | 4.216% | 6.767% |  | 4.187% | 0% | 1.293% | 4.556% | 10.1% | 3.098% | 0% | 5.314% | 3.355%
| PDM3 |  | 3.242% | -2.124% | 0.7838% | 10.62% | 14.3% |  | -9.049% | -26.26% | -1.797% | -16.42% | -5.146% | 22.91% | -11.51% | 4.471% | -10.28%
| Z-Score (Absolute) |  | 0.28 | 0.12 | 0.1 | 0.82 | 1.09 | 264.21 | 0.63 | NA | 0.1 | 1.17 | 0.34 | 1.72 | 0.81 | 0.37 | 0.72
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 250c (Indicative Value 0.11 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | [16 empty] | 2
|  |  | 25g | 0.2g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.118 | 0.11201 | 0.1 | [16 empty] | 1
| 1 | 2 | 0.112 | 0.10008 | 0.1 | [16 empty] | 7
| 1 | 3 | 0.114 | 0.10439 | 0.1 | [16 empty] | 16
| 1 | 4 | 0.114 | 0.10597 | 0.1 | [16 empty] | 0.10676
| 1 | 5 | 0.114 | 0.10676 | 0.1 | [16 empty] | 13
| 1 | 6 | 0.114 | 0.10647 | 0.1
| Mean |  | 0.11433 | 0.10595 | 0.1
| Median |  | 0.114 | 0.10622 | 0.1
| Std Dev. |  | 0.0019664 | 0.0038603 | 1.5202e-17
| Rel.Std.Dev. |  | 1.72% | 3.644% | 1.52e-14%
| PDM3 |  | 7.094% | -0.7614% | -6.332%
| Z-Score (Absolute) |  | 0.95 | 0 | 0.67
| Analytical results for U in OREAS 250c (Certified Value 0.97 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 1 | 0.93642 | 0.97 | 1 | 0.89 | 1 | 0.8 | 1.034 | 0.82 | 0.95194 | 1.06 | 0.91 | 1.04 | 0.87 | 1.07 |  |  |  |  | 1
| 1 | 2 | 1 | 0.87883 | 0.98 | 1 | 0.89 | 0.96 | 1 | 1.033 | 0.81 | 1.0095 | 1.14 | 0.87 | 1.09 | 0.9 | 0.99 |  |  |  |  | 30
| 1 | 3 | 1 | 0.90135 | 0.94 | 1 | 0.9 | 0.92 | 1 | 1.034 | 0.82 | 1.0237 | 1.09 | 1.01 | 1.04 | 0.89 | 1.03 |  |  |  |  | 16
| 1 | 4 | 1 | 0.91971 | 0.98 | 0.9 | 0.9 | 0.96 | 1 | 1.02 | 0.89 | 1.0548 | 1.12 | 0.91 | 1.07 | 0.93 | 1 |  |  |  |  | 0.96523
| 1 | 5 | 1.02 | 0.91016 | 0.95 | 1 | 0.89 | 0.91 | 1 | 1.039 | 0.81 | 1.0031 | 1.12 | 0.89 | 1.02 | 0.94 | 1.06 |  |  |  |  | 114
| 1 | 6 | 1.02 | 0.95456 | 0.96 | 1 | 0.89 | 0.98 | 0.9 | 0.985 | 0.8 | 0.9497 | 1.08 | 0.92 | 1 | 0.88 | 0.99
| Mean |  | 1.0067 | 0.91684 | 0.96333 | 0.98333 | 0.89333 | 0.955 | 0.95 | 1.0242 | 0.825 | 0.99879 | 1.1017 | 0.91833 | 1.0433 | 0.90167 | 1.0233
| Median |  | 1 | 0.91493 | 0.965 | 1 | 0.89 | 0.96 | 1 | 1.0335 | 0.815 | 1.0063 | 1.105 | 0.91 | 1.04 | 0.895 | 1.015
| Std Dev. |  | 0.010328 | 0.026622 | 0.01633 | 0.040825 | 0.005164 | 0.034496 | 0.083666 | 0.020213 | 0.032711 | 0.04122 | 0.029944 | 0.048339 | 0.03266 | 0.027869 | 0.03559
| Rel.Std.Dev. |  | 1.026% | 2.904% | 1.695% | 4.152% | 0.5781% | 3.612% | 8.807% | 1.974% | 3.965% | 4.127% | 2.718% | 5.264% | 3.13% | 3.091% | 3.478%
| PDM3 |  | 4.293% | -5.013% | -0.1962% | 1.876% | -7.448% | -1.06% | -1.578% | 6.106% | -14.53% | 3.477% | 14.14% | -4.858% | 8.092% | -6.585% | 6.02%
| Z-Score (Absolute) |  | 0.49 | 0.52 | 0 | NA | 0.79 | 0.09 | NA | 0.68 | 1.55 | 0.4 | 1.55 | 0.51 | 0.9 | 0.69 | 0.67
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 250c (Certified Value 33.7 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 34 | 34.4771 | 31.709 | 33 | 36 | 30 | 31 | 36.245 | 33 | 37 | 31 | 34.1875 | 35 | 35 | 33 | 32 |  |  |  | 1
| 1 | 2 | 34 | 34.1144 | 33.025 | 34 | 37 | 32 | 30 | 35.6832 | 33 | 35 | 32 | 36.2222 | 36 | 34 | 33 | 32 |  |  |  | 31
| 1 | 3 | 35 | 34.4459 | 33.453 | 36 | 36 | 31 | 31 | 36.2533 | 33 | 34 | 32 | 35.5595 | 34 | 34 | 32 | 32 |  |  |  | 16
| 1 | 4 | 36 | 35.1522 | 32.92 | 36 | 36 | 30 | 32 | 35.7935 | 34 | 34 | 31 | 34.7595 | 35 | 35 | 33 | 32 |  |  |  | 33.6989
| 1 | 5 | 35 | 34.4937 | 32.403 | 36 | 35 | 31 | 31 | 35.7019 | 33 | 36 | 31 | 35.6033 | 34 | 33 | 32 | 32 |  |  |  | 115
| 1 | 6 | 34 | 34.82 | 32.874 | 37 | 36 | 33 | 31 | 35.6856 | 34 | 32 | 31 | 33.5127 | 35 | 34 | 32 | 32
| Mean |  | 34.6667 | 34.5839 | 32.7307 | 35.3333 | 36 | 31.1667 | 31 | 35.8938 | 33.3333 | 34.6667 | 31.3333 | 34.9741 | 34.8333 | 34.1667 | 32.5 | 32
| Median |  | 34.5 | 34.4854 | 32.897 | 36 | 36 | 31 | 31 | 35.7477 | 33 | 34.5 | 31 | 35.1595 | 35 | 34 | 32.5 | 32
| Std Dev. |  | 0.8165 | 0.35711 | 0.60265 | 1.5055 | 0.63246 | 1.169 | 0.63246 | 0.27827 | 0.5164 | 1.7512 | 0.5164 | 1.0096 | 0.75277 | 0.75277 | 0.54772 | 0
| Rel.Std.Dev. |  | 2.355% | 1.033% | 1.841% | 4.261% | 1.757% | 3.751% | 2.04% | 0.7752% | 1.549% | 5.052% | 1.648% | 2.887% | 2.161% | 2.203% | 1.685% | 0%
| PDM3 |  | 2.872% | 2.626% | -2.873% | 4.85% | 6.828% | -7.514% | -8.009% | 6.513% | -1.085% | 2.872% | -7.02% | 3.784% | 3.366% | 1.388% | -3.558% | -5.041%
| Z-Score (Absolute) |  | 0.15 | 0.11 | 0.87 | 0.5 | 0.86 | 1.69 | 1.78 | 0.8 | 0.55 | 0.15 | 1.6 | 0.32 | 0.24 | 0.11 | 0.99 | 1.25
| Analytical results for W in OREAS 250c (Certified Value 0.38 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.37 | < 1 | 2 | 0.2 | 0.3 | 0.2 | < 5 | 0.39 | 0.2 | 0.42 | 0.33971 | 0.3 | 0.38 | 0.46 | 0.32 | 0.27 |  |  |  | 1
| 1 | 2 | 0.39 | < 1 | 1 | 0.3 | 0.3 | 0.18 | < 5 | 0.39 | 0.3 | 0.37 | 0.33971 | 0.3 | 0.37 | 0.42 | 0.32 | 0.29 |  |  |  | 32
| 1 | 3 | 0.36 | < 1 | < 1 | 0.3 | 0.3 | 0.19 | < 5 | 0.37 | 0.3 | 0.42 | 0.38431 | 0.4 | 0.42 | 0.44 | 0.34 | 0.29 |  |  |  | 16
| 1 | 4 | 0.39 | < 1 | 1 | 0.3 | 0.3 | 0.18 | < 5 | 0.4 | 0.3 | 0.39 | 0.35311 | 0.3 | 0.36 | 0.45 | 0.32 | 0.26 |  |  |  | 0.38335
| 1 | 5 | 0.37 | < 1 | 1 | 0.3 | 0.3 | 0.2 | < 5 | 0.38 | 0.3 | 0.39 | 0.37843 | 0.3 | 0.4 | 0.48 | 0.32 | 0.31 |  |  |  | 116
| 1 | 6 | 0.4 | < 1 | < 1 | 0.2 | 0.3 | 0.18 | < 5 | 0.41 | 0.3 | 0.39 | 0.38549 | 0.3 | 0.38 | 0.42 | 0.32 | 0.32
| Mean |  | 0.38 |  | 1.25 | 0.26667 | 0.3 | 0.18833 |  | 0.39 | 0.28333 | 0.39667 | 0.36346 | 0.31667 | 0.385 | 0.445 | 0.32333 | 0.29
| Median |  | 0.38 |  | 1 | 0.3 | 0.3 | 0.185 |  | 0.39 | 0.3 | 0.39 | 0.36577 | 0.3 | 0.38 | 0.445 | 0.32 | 0.29
| Std Dev. |  | 0.015492 |  | 0.5 | 0.05164 | 0 | 0.0098319 |  | 0.014142 | 0.040825 | 0.019664 | 0.021817 | 0.040825 | 0.021679 | 0.023452 | 0.008165 | 0.022804
| Rel.Std.Dev. |  | 4.077% |  | 40% | 19.36% | 0% | 5.22% |  | 3.626% | 14.41% | 4.957% | 6.003% | 12.89% | 5.631% | 5.27% | 2.525% | 7.863%
| PDM3 |  | -0.8743% |  | 226.1% | -30.44% | -21.74% | -50.87% |  | 1.734% | -26.09% | 3.473% | -5.189% | -17.4% | 0.43% | 16.08% | -15.66% | -24.35%
| Z-Score (Absolute) |  | 0.06 | 1.29 | NA | NA | NA | 2.21 | 23.77 | 0.06 | NA | 0.13 | 0.24 | NA | 0 | 0.67 | 0.69 | 1.07
| Analytical results for Y in OREAS 250c (Certified Value 9.1 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 9.32 | 9.4029 | 10 | 8.78 | 8.95 | 8.38 | 9.09 | 8.87 | 9.25 | 8.8629 | 9.1 | 9.54 | 9.94 | 9.1 | 8.83 |  |  |  |  | 1
| 1 | 2 | 8.93 | 9.2644 | 10 | 9.13 | 9.12 | 8.3 | 9.1533 | 9 | 9.16 | 9.3347 | 9.3 | 9.32 | 10.45 | 8.98 | 8.82 |  |  |  |  | 33
| 1 | 3 | 9.35 | 9.3703 | 11 | 8.76 | 8.73 | 8.32 | 9.2833 | 8.81 | 9.46 | 9.254 | 9.1 | 9.52 | 9.89 | 9.23 | 8.69 |  |  |  |  | 16
| 1 | 4 | 9.49 | 9.5841 | 11 | 8.85 | 8.23 | 8.37 | 9.2167 | 8.88 | 9.25 | 9.1324 | 9 | 9.45 | 10 | 9.34 | 8.73 |  |  |  |  | 9.1007
| 1 | 5 | 9.23 | 9.3223 | 11 | 8.75 | 8.73 | 8.3 | 9.1467 | 8.61 | 9.34 | 9.3427 | 9.1 | 9.12 | 9.74 | 9.16 | 8.76 |  |  |  |  | 117
| 1 | 6 | 9.48 | 9.5088 | 11 | 8.47 | 9.07 | 8.29 | 9.1567 | 9.04 | 9.4 | 8.7788 | 9.2 | 9.23 | 9.84 | 9.22 | 8.72
| Mean |  | 9.3 | 9.4088 | 10.6667 | 8.79 | 8.805 | 8.3267 | 9.1744 | 8.8683 | 9.31 | 9.1176 | 9.1333 | 9.3633 | 9.9767 | 9.1717 | 8.7583
| Median |  | 9.335 | 9.3866 | 11 | 8.77 | 8.84 | 8.31 | 9.155 | 8.875 | 9.295 | 9.1932 | 9.1 | 9.385 | 9.915 | 9.19 | 8.745
| Std Dev. |  | 0.20649 | 0.11881 | 0.5164 | 0.21175 | 0.3263 | 0.038816 | 0.066789 | 0.15303 | 0.11063 | 0.24341 | 0.10328 | 0.1686 | 0.24825 | 0.12336 | 0.056362
| Rel.Std.Dev. |  | 2.22% | 1.263% | 4.841% | 2.409% | 3.706% | 0.4662% | 0.728% | 1.726% | 1.188% | 2.67% | 1.131% | 1.801% | 2.488% | 1.345% | 0.6435%
| PDM3 |  | 2.19% | 3.386% | 17.21% | -3.414% | -3.249% | -8.505% | 0.8105% | -2.553% | 2.3% | 0.1859% | 0.3588% | 2.886% | 9.626% | 0.78% | -3.762%
| Z-Score (Absolute) |  | 0.43 | 0.74 | NA | 1.05 | 1 | 2.38 | 0.06 | 0.82 | 0.45 | 0.1 | 0.06 | 0.61 | 2.38 | 0.06 | 1.14
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 250c (Certified Value 0.63 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.17 | 1.22 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.25g | [14 empty] | 3
| 1 | 1 | 0.7 | 0.64632 | 0.7 | 0.6 | 0.6 | [14 empty] | 1
| 1 | 2 | 0.672 | 0.6012 | 0.6 | 0.6 | 0.6 | [14 empty] | 34
| 1 | 3 | 0.669 | 0.63443 | 0.6 | 0.6 | 0.6 | [14 empty] | 16
| 1 | 4 | 0.695 | 0.66014 | 0.6 | 0.6 | 0.6 | [14 empty] | 0.63153
| 1 | 5 | 0.687 | 0.64863 | 0.6 | 0.6 | 0.6 | [14 empty] | 118
| 1 | 6 | 0.701 | 0.63131 | 0.7 | 0.6 | 0.6
| Mean |  | 0.68733 | 0.63701 | 0.63333 | 0.6 | 0.6
| Median |  | 0.691 | 0.64038 | 0.6 | 0.6 | 0.6
| Std Dev. |  | 0.013981 | 0.020388 | 0.05164 | 0 | 0
| Rel.Std.Dev. |  | 2.034% | 3.201% | 8.154% | 0% | 0%
| PDM3 |  | 8.835% | 0.8664% | 0.2848% | -4.993% | -4.993%
| Z-Score (Absolute) |  | 1.09 | 0.07 | 0 | 0.67 | 0.67
| Analytical results for Zn in OREAS 250c (Certified Value 46.4 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 47 | 55.4796 | 47.5159 | 36.354 | 41 | 50 | 53.5 | 44 | 40.2233 | 47 | 48 | 41.78 | 44 | 47.2314 | 52.3 | 49 | 45 | 44 | 40 | 1
| 1 | 2 | 46 | 54.4968 | 47.044 | 38.089 | 44 | 50 | 53.2 | 44 | 40.5933 | 48 | 47 | 41.325 | 44 | 52.8252 | 53.8 | 48 | 45 | 44 | 40 | 35
| 1 | 3 | 47 | 56.7216 | 47.7464 | 38.64 | 46 | 51 | 52.8 | 43 | 41.7067 | 48 | 46 | 41.201 | 44 | 50.3169 | 52.5 | 49 | 44 | 44 | 40 | 16
| 1 | 4 | 48 | 55.2096 | 48.1998 | 38.309 | 46 | 50 | 53.3 | 45 | 40.81 | 48 | 46 | 41.722 | 43 | 50.6526 | 52 | 49 | 45 | 44 | 40 | 46.3867
| 1 | 5 | 47 | 55.8036 | 47.4192 | 37.073 | 47 | 51 | 53.9 | 46 | 40.41 | 49 | 48 | 41.429 | 44 | 50.7356 | 52.7 | 47 | 45 | 44 | 40 | 119
| 1 | 6 | 48 | 53.2116 | 47.834 | 35.829 | 46 | 51 | 53.6 | 44 | 41.0367 | 48 | 43 | 41.491 | 44 | 46.6179 | 52.6 | 48 | 45 | 44 | 40
| Mean |  | 47.1667 | 55.1538 | 47.6266 | 37.3823 | 45 | 50.5 | 53.3833 | 44.3333 | 40.7967 | 48 | 46.3333 | 41.4913 | 43.8333 | 49.7299 | 52.65 | 48.3333 | 44.8333 | 44 | 40
| Median |  | 47 | 55.3446 | 47.6312 | 37.581 | 46 | 50.5 | 53.4 | 44 | 40.7017 | 48 | 46.5 | 41.46 | 44 | 50.4848 | 52.55 | 48.5 | 45 | 44 | 40
| Std Dev. |  | 0.75277 | 1.1994 | 0.39514 | 1.1407 | 2.1909 | 0.54772 | 0.37639 | 1.0328 | 0.53018 | 0.63246 | 1.8619 | 0.22467 | 0.40825 | 2.3544 | 0.61563 | 0.8165 | 0.40825 | 0 | 0
| Rel.Std.Dev. |  | 1.596% | 2.175% | 0.8297% | 3.051% | 4.869% | 1.085% | 0.7051% | 2.33% | 1.3% | 1.318% | 4.018% | 0.5415% | 0.9314% | 4.734% | 1.169% | 1.689% | 0.9106% | 0% | 0%
| PDM3 |  | 1.681% | 18.9% | 2.673% | -19.41% | -2.989% | 8.867% | 15.08% | -4.427% | -12.05% | 3.478% | -0.115% | -10.55% | -5.505% | 7.207% | 13.5% | 4.197% | -3.349% | -5.145% | -13.77%
| Z-Score (Absolute) |  | 0.22 | 2.38 | 0.35 | 2.41 | 0.36 | 1.12 | 1.9 | 0.54 | 1.49 | 0.45 | 0 | 1.31 | 0.67 | 0.92 | 1.7 | 0.54 | 0.4 | 0.63 | 1.71
| Analytical results for Zr in OREAS 250c (Certified Value 37.1 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 20.9 | 47.0002 | 88 | 42.4 | 4.3 | 35 | 31.7381 | 35 | 35.6 | 49.5992 | 47.7 | 25.1 | 34.9 | 25.1 | 49.8 |  |  |  |  | 1
| 1 | 2 | 20.8 | 46.0248 | 93 | 42.9 | 5.9 | 35 | 32.213 | 27.1 | 28.7 | 57.7533 | 49.5 | 23.3 | 37.1 | 21.9 | 49 |  |  |  |  | 36
| 1 | 3 | 21.1 | 47.3929 | 94 | 41.6 | 7.6 | 34 | 32.8459 | 38.2 | 38.4 | 58.1728 | 49.3 | 35.8 | 35.7 | 21.4 | 49.8 |  |  |  |  | 16
| 1 | 4 | 19.7 | 46.7274 | 130 | 43.2 | 4.9 | 37 | 32.5084 | 34.1 | 31.5 | 57.1444 | 46 | 18.2 | 36.4 | 26.6 | 50.9 |  |  |  |  | 37.1153
| 1 | 5 | 17.4 | 46.8612 | 93 | 42 | 4.8 | 36 | 31.7992 | 28 | 34.1 | 50.8367 | 42.8 | 30.6 | 43.6 | 20.8 | 49.6 |  |  |  |  | 120
| 1 | 6 | 18.2 | 46.0247 | 96 | 41.1 | 4.1 | 34 | 33.1742 | 36.8 | 38 | 53.9026 | 44.8 | 28.9 | 41.1 | 20.9 | 49
| Mean |  | 19.6833 | 46.6718 | 99 | 42.2 | 5.2667 | 35.1667 | 32.3798 | 33.2 | 34.3833 | 54.5682 | 46.6833 | 26.9833 | 38.1333 | 22.7833 | 49.6833
| Median |  | 20.25 | 46.7943 | 93.5 | 42.2 | 4.85 | 35 | 32.3607 | 34.55 | 34.85 | 55.5235 | 46.85 | 27 | 36.75 | 21.65 | 49.7
| Std Dev. |  | 1.5587 | 0.54849 | 15.4143 | 0.79246 | 1.3033 | 1.169 | 0.573 | 4.61 | 3.7786 | 3.7103 | 2.6423 | 6.1441 | 3.439 | 2.4539 | 0.69976
| Rel.Std.Dev. |  | 7.919% | 1.175% | 15.57% | 1.878% | 24.75% | 3.324% | 1.77% | 13.89% | 10.99% | 6.799% | 5.66% | 22.77% | 9.018% | 10.77% | 1.408%
| PDM3 |  | -46.97% | 25.75% | 166.7% | 13.7% | -85.81% | -5.25% | -12.76% | -10.55% | -7.361% | 47.02% | 25.78% | -27.3% | 2.743% | -38.61% | 33.86%
| Z-Score (Absolute) |  | 0.91 | 0.67 | 3.74 | 0.41 | 1.75 | 0 | 0.16 | 0.12 | 0.05 | 1.14 | 0.67 | 0.48 | 0.17 | 0.73 | 0.85
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 250c (Indicative Value 10.07 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.06 | 1
| 1 | 2 | 10.07 | 11
| Mean |  | 10.065 | 16
| Median |  | 10.065 | 10.065
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.07025%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 250c (Indicative Value 2.35 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.34 | 1
| 1 | 2 | 2.36 | 12
| Mean |  | 2.35 | 16
| Median |  | 2.35 | 2.35
| Std Dev. |  | 0.014142 | 18
| Rel.Std.Dev. |  | 0.6018%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 250c (Indicative Value 5.9 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.9 | 1
| 1 | 2 | 5.9 | 13
| Mean |  | 5.9 | 16
| Median |  | 5.9 | 5.9
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 250c (Indicative Value 1.58 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.58 | 1
| 1 | 2 | 1.58 | 14
| Mean |  | 1.58 | 16
| Median |  | 1.58 | 1.58
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 250c (Indicative Value 2.67 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.66 | 1
| 1 | 2 | 2.67 | 11
| Mean |  | 2.665 | 16
| Median |  | 2.665 | 2.665
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.2653%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 250c (Indicative Value 0.064 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.064 | 1
| 1 | 2 | 0.064 | 12
| Mean |  | 0.064 | 16
| Median |  | 0.064 | 0.064
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 250c (Indicative Value 1.46 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.46 | 1
| 1 | 2 | 1.46 | 13
| Mean |  | 1.46 | 16
| Median |  | 1.46 | 1.46
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 250c (Indicative Value 0.219 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.218 | 1
| 1 | 2 | 0.219 | 14
| Mean |  | 0.2185 | 16
| Median |  | 0.2185 | 0.2185
| Std Dev. |  | 0.00070711 | 20
| Rel.Std.Dev. |  | 0.3236%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 250c (Indicative Value 72.25 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 72.25 | 1
| 1 | 2 | 72.24 | 11
| Mean |  | 72.245 | 16
| Median |  | 72.245 | 72.245
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.009788%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 250c (Indicative Value 0.07 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.069 | 12
| Mean |  | 0.0695 | 16
| Median |  | 0.0695 | 0.0695
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 1.017%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 250c (Indicative Value 0.997 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.996 | 1
| 1 | 2 | 0.998 | 13
| Mean |  | 0.997 | 16
| Median |  | 0.997 | 0.997
| Std Dev. |  | 0.0014142 | 19
| Rel.Std.Dev. |  | 0.1418%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 250c (Indicative Value 2.16 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.15 | 1
| 1 | 2 | 2.16 | 16
| Mean |  | 2.155 | 16
| Median |  | 2.155 | 2.155
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 0.3281%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 250c (Indicative Value 0.075 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.08 | 18
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 250c (Indicative Value 0.01 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | 0.01 | 18
| Mean |  | 0.01 | 16
| Median |  | 0.01 | 0.01
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 250c (Indicative Value 0.1 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.1 | 20
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.1
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 250c (Indicative Value 51 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 50.6 | 1
| 1 | 2 | 50.6 | 21
| Mean |  | 50.6 | 16
| Median |  | 50.6 | 50.6
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 250c (Indicative Value 350 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 345 | 1
| 1 | 2 | 354 | 22
| Mean |  | 349.5 | 16
| Median |  | 349.5 | 349.5
| Std Dev. |  | 6.364 | 28
| Rel.Std.Dev. |  | 1.821%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 250c (Indicative Value 2 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 2.2 | 23
| Mean |  | 2 | 16
| Median |  | 2 | 2
| Std Dev. |  | 0.28284 | 29
| Rel.Std.Dev. |  | 14.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 250c (Indicative Value 0.3 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.28 | 1
| 1 | 2 | 0.32 | 24
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0.028284 | 30
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 250c (Indicative Value < 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 25
| Mean |  |  | 16
| Median |  |  | < 0.1
| Std Dev. |  |  | 31
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 250c (Indicative Value 68 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 67.8 | 1
| 1 | 2 | 67.9 | 26
| Mean |  | 67.85 | 16
| Median |  | 67.85 | 67.85
| Std Dev. |  | 0.070711 | 32
| Rel.Std.Dev. |  | 0.1042%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 250c (Indicative Value 20.5 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.4 | 1
| 1 | 2 | 20.6 | 27
| Mean |  | 20.5 | 16
| Median |  | 20.5 | 20.5
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 250c (Indicative Value 115 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 114 | 1
| 1 | 2 | 116 | 28
| Mean |  | 115 | 16
| Median |  | 115 | 115
| Std Dev. |  | 1.4142 | 34
| Rel.Std.Dev. |  | 1.23%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 250c (Indicative Value 2.52 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.48 | 1
| 1 | 2 | 2.56 | 29
| Mean |  | 2.52 | 16
| Median |  | 2.52 | 2.52
| Std Dev. |  | 0.056569 | 35
| Rel.Std.Dev. |  | 2.245%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 250c (Indicative Value 32 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36 | 1
| 1 | 2 | 28 | 30
| Mean |  | 32 | 16
| Median |  | 32 | 32
| Std Dev. |  | 5.6569 | 36
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 250c (Indicative Value 4.12 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.16 | 1
| 1 | 2 | 4.08 | 4
| Mean |  | 4.12 | 16
| Median |  | 4.12 | 4.12
| Std Dev. |  | 0.056569 | 37
| Rel.Std.Dev. |  | 1.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 250c (Indicative Value 2.04 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.96 | 1
| 1 | 2 | 2.11 | 5
| Mean |  | 2.035 | 16
| Median |  | 2.035 | 2.035
| Std Dev. |  | 0.10607 | 38
| Rel.Std.Dev. |  | 5.212%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 250c (Indicative Value 1.59 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.58 | 1
| 1 | 2 | 1.6 | 6
| Mean |  | 1.59 | 16
| Median |  | 1.59 | 1.59
| Std Dev. |  | 0.014142 | 39
| Rel.Std.Dev. |  | 0.8894%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 250c (Indicative Value 14.6 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.5 | 1
| 1 | 2 | 14.6 | 34
| Mean |  | 14.55 | 16
| Median |  | 14.55 | 14.55
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 0.486%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 250c (Indicative Value 5.33 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.19 | 1
| 1 | 2 | 5.47 | 7
| Mean |  | 5.33 | 16
| Median |  | 5.33 | 5.33
| Std Dev. |  | 0.19799 | 41
| Rel.Std.Dev. |  | 3.715%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 250c (Indicative Value 1.08 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.1 | 1
| 1 | 2 | 1.05 | 1
| Mean |  | 1.075 | 16
| Median |  | 1.075 | 1.075
| Std Dev. |  | 0.035355 | 42
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 250c (Indicative Value 5.99 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6 | 1
| 1 | 2 | 5.97 | 8
| Mean |  | 5.985 | 16
| Median |  | 5.985 | 5.985
| Std Dev. |  | 0.021213 | 26
| Rel.Std.Dev. |  | 0.3544%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 250c (Indicative Value 0.77 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.78 | 4
| Mean |  | 0.77 | 16
| Median |  | 0.77 | 0.77
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 1.837%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 250c (Indicative Value 0.038 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | < 0.05 | 22
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.0375
| Std Dev. |  |  | 28
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 250c (Indicative Value 38.3 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.2 | 1
| 1 | 2 | 38.4 | 23
| Mean |  | 38.3 | 16
| Median |  | 38.3 | 38.3
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 0.3692%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 250c (Indicative Value 0.23 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.24 | 5
| Mean |  | 0.23 | 16
| Median |  | 0.23 | 0.23
| Std Dev. |  | 0.014142 | 30
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 250c (Indicative Value 0.051 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0503 | 1
| 1 | 2 | 0.0514 | 25
| Mean |  | 0.05085 | 16
| Median |  | 0.05085 | 0.05085
| Std Dev. |  | 0.00077782 | 31
| Rel.Std.Dev. |  | 1.53%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 250c (Indicative Value 1.8 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.8 | 26
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 250c (Indicative Value 32.6 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32.4 | 1
| 1 | 2 | 32.8 | 27
| Mean |  | 32.6 | 16
| Median |  | 32.6 | 32.6
| Std Dev. |  | 0.28284 | 33
| Rel.Std.Dev. |  | 0.8676%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 250c (Indicative Value 33.1 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32.9 | 1
| 1 | 2 | 33.2 | 6
| Mean |  | 33.05 | 16
| Median |  | 33.05 | 33.05
| Std Dev. |  | 0.21213 | 34
| Rel.Std.Dev. |  | 0.6419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 250c (Indicative Value 86 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84 | 1
| 1 | 2 | 88 | 29
| Mean |  | 86 | 16
| Median |  | 86 | 86
| Std Dev. |  | 2.8284 | 35
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 250c (Indicative Value 10.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11 | 1
| 1 | 2 | 10 | 30
| Mean |  | 10.5 | 16
| Median |  | 10.5 | 10.5
| Std Dev. |  | 0.70711 | 36
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 250c (Indicative Value 8.9 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.92 | 1
| 1 | 2 | 8.88 | 8
| Mean |  | 8.9 | 16
| Median |  | 8.9 | 8.9
| Std Dev. |  | 0.028284 | 37
| Rel.Std.Dev. |  | 0.3178%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 250c (Indicative Value 63 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 62.9 | 1
| 1 | 2 | 62.8 | 32
| Mean |  | 62.85 | 16
| Median |  | 62.85 | 62.85
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.1125%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 250c (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 33
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 39
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 250c (Indicative Value 0.9 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.9 | 1
| 1 | 2 | 0.9 | 34
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 250c (Indicative Value 9.55 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.7 | 1
| 1 | 2 | 9.4 | 35
| Mean |  | 9.55 | 16
| Median |  | 9.55 | 9.55
| Std Dev. |  | 0.21213 | 41
| Rel.Std.Dev. |  | 2.221%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 250c (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 1
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 42
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 250c (Indicative Value 6.32 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.52 | 1
| 1 | 2 | 6.11 | 5
| Mean |  | 6.315 | 16
| Median |  | 6.315 | 6.315
| Std Dev. |  | 0.28991 | 26
| Rel.Std.Dev. |  | 4.591%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 250c (Indicative Value 4.5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.6 | 21
| Mean |  | 4.5 | 16
| Median |  | 4.5 | 4.5
| Std Dev. |  | 0.14142 | 27
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 250c (Indicative Value 296 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 294 | 1
| 1 | 2 | 298 | 22
| Mean |  | 296 | 16
| Median |  | 296 | 296
| Std Dev. |  | 2.8284 | 28
| Rel.Std.Dev. |  | 0.9555%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 250c (Indicative Value 2.21 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.21 | 1
| 1 | 2 | 2.2 | 23
| Mean |  | 2.205 | 16
| Median |  | 2.205 | 2.205
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.3207%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 250c (Indicative Value 0.77 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.77 | 1
| 1 | 2 | 0.77 | 24
| Mean |  | 0.77 | 16
| Median |  | 0.77 | 0.77
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 250c (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 2
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 31
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 250c (Indicative Value 9.09 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.96 | 1
| 1 | 2 | 9.22 | 26
| Mean |  | 9.09 | 16
| Median |  | 9.09 | 9.09
| Std Dev. |  | 0.18385 | 32
| Rel.Std.Dev. |  | 2.023%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 250c (Indicative Value 0.599 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.597 | 27
| Mean |  | 0.5985 | 16
| Median |  | 0.5985 | 0.5985
| Std Dev. |  | 0.0021213 | 33
| Rel.Std.Dev. |  | 0.3544%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 250c (Indicative Value 0.15 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | 0.2 | 28
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 34
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 250c (Indicative Value 0.29 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.28 | 1
| 1 | 2 | 0.29 | 7
| Mean |  | 0.285 | 16
| Median |  | 0.285 | 0.285
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 250c (Indicative Value 1.81 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.78 | 1
| 1 | 2 | 1.84 | 30
| Mean |  | 1.81 | 16
| Median |  | 1.81 | 1.81
| Std Dev. |  | 0.042426 | 36
| Rel.Std.Dev. |  | 2.344%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 250c (Indicative Value 73 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72.9 | 1
| 1 | 2 | 72.8 | 31
| Mean |  | 72.85 | 16
| Median |  | 72.85 | 72.85
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.09706%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 250c (Indicative Value 2.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.5 | 1
| 1 | 2 | 2.5 | 32
| Mean |  | 2.5 | 16
| Median |  | 2.5 | 2.5
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 250c (Indicative Value 19.8 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.5 | 1
| 1 | 2 | 20 | 33
| Mean |  | 19.75 | 16
| Median |  | 19.75 | 19.75
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 250c (Indicative Value 1.75 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.75 | 1
| 1 | 2 | 1.74 | 34
| Mean |  | 1.745 | 16
| Median |  | 1.745 | 1.745
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.4052%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 250c (Indicative Value 75 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75 | 1
| 1 | 2 | 75 | 35
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 250c (Indicative Value 243 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 241 | 1
| 1 | 2 | 245 | 36
| Mean |  | 243 | 16
| Median |  | 243 | 243
| Std Dev. |  | 2.8284 | 42
| Rel.Std.Dev. |  | 1.164%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.