## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 263b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.21447 | 0.20984 | 0.2191 | 0.21223 | 0.21672
PhotonAssay (recommended gross mass*490-520 g)
Au, Gold (ppm) | 0.21103 | 0.20483 | 0.21724 | 0.21039 | 0.21168
Aqua Regia Digestion (sample mass 10-50g)
Au, Gold (ppm) | 0.19108 | 0.18481 | 0.19734 | 0.18889 | 0.19326
Cyanide Leach
Au, Gold (ppm) | 0.19098 | 0.18522 | 0.19675 | 0.19021 | 0.19175
4-Acid Digestion
Ag, Silver (ppm) | 0.30715 | 0.27738 | 0.33691 | 0.28058 | 0.33372
Al, Aluminium (wt.%) | 6.2983 | 6.151 | 6.4456 | 6.1755 | 6.4211
As, Arsenic (ppm) | 257.4637 | 250.8738 | 264.0536 | 252.4177 | 262.5097
Ba, Barium (ppm) | 479.1494 | 466.4899 | 491.8088 | 471.2083 | 487.0904
Be, Beryllium (ppm) | 2.1588 | 2.0567 | 2.261 | 2.0941 | 2.2236
Bi, Bismuth (ppm) | 10.1057 | 9.3439 | 10.8675 | 9.6042 | 10.6073
Ca, Calcium (wt.%) | 0.29448 | 0.28417 | 0.3048 | 0.28662 | 0.30235
Cd, Cadmium (ppm) | 0.04369 | 0.023663 | 0.063717 | IND | IND
Ce, Cerium (ppm) | 82.2028 | 78.9554 | 85.4503 | 80.2765 | 84.1291
Co, Cobalt (ppm) | 4.8092 | 4.4792 | 5.1391 | 4.6529 | 4.9654
Cr, Chromium (ppm) | 78.1309 | 74.0155 | 82.2464 | 75.7497 | 80.5122
Cs, Caesium (ppm) | 8.0141 | 7.7406 | 8.2876 | 7.839 | 8.1892
Cu, Copper (ppm) | 74.9709 | 71.1183 | 78.8234 | 72.5055 | 77.4362
Dy, Dysprosium (ppm) | 3.2126 | 2.9539 | 3.4712 | 3.0831 | 3.3421
Er, Erbium (ppm) | 1.6664 | 1.474 | 1.8588 | 1.5805 | 1.7523
Eu, Europium (ppm) | 1.0244 | 0.93427 | 1.1145 | 0.98575 | 1.063
Fe, Iron (wt.%) | 2.6821 | 2.6171 | 2.7472 | 2.6346 | 2.7297
Ga, Gallium (ppm) | 16.5685 | 16.0006 | 17.1365 | 16.1505 | 16.9865
Gd, Gadolinium (ppm) | 4.2644 | 3.9047 | 4.6241 | 4.0439 | 4.4849
Ge, Germanium (ppm) | 0.11739 | 0.092872 | 0.14191 | 0.089003 | 0.14577
Hf, Hafnium (ppm) | 3.3742 | 3.1869 | 3.5614 | 3.1829 | 3.5655
Ho, Holmium (ppm) | 0.57601 | 0.5156 | 0.63642 | 0.55366 | 0.59836
In, Indium (ppm) | 0.080845 | 0.070136 | 0.091554 | 0.075151 | 0.086539
K, Potassium (wt.%) | 2.0324 | 1.9822 | 2.0825 | 1.9897 | 2.075
La, Lanthanum (ppm) | 40.2916 | 38.3962 | 42.1871 | 39.1636 | 41.4197
Li, Lithium (ppm) | 37.9046 | 36.7685 | 39.0407 | 37.2132 | 38.5959
Lu, Lutetium (ppm) | 0.2664 | 0.22722 | 0.30558 | 0.22962 | 0.30318
Mg, Magnesium (wt.%) | 0.52699 | 0.50807 | 0.54591 | 0.51789 | 0.53609
Mn, Manganese (wt.%) | 0.017356 | 0.016772 | 0.017941 | 0.016942 | 0.01777
Mo, Molybdenum (ppm) | 2.3601 | 2.211 | 2.5092 | 2.2357 | 2.4844
Na, Sodium (wt.%) | 0.11487 | 0.1098 | 0.11994 | 0.11114 | 0.1186
Nb, Niobium (ppm) | 10.0686 | 9.1927 | 10.9446 | 9.7034 | 10.4338
Nd, Neodymium (ppm) | 33.0779 | 31.3644 | 34.7913 | 32.3366 | 33.8191
Ni, Nickel (ppm) | 44.0176 | 42.6408 | 45.3945 | 42.6849 | 45.3503
P, Phosphorus (wt.%) | 0.029386 | 0.028396 | 0.030377 | 0.028399 | 0.030374
Pb, Lead (ppm) | 19.3605 | 18.5829 | 20.138 | 18.7666 | 19.9543
Pr, Praseodymium (ppm) | 9.184 | 8.6029 | 9.7652 | 8.9012 | 9.4669
Rb, Rubidium (ppm) | 131.5962 | 126.97 | 136.2224 | 128.4507 | 134.7417
S, Sulphur (wt.%) | 0.028769 | 0.025631 | 0.031907 | 0.027346 | 0.030192
Sb, Antimony (ppm) | 11.3175 | 10.4864 | 12.1485 | 10.937 | 11.6979
Sc, Scandium (ppm) | 10.9153 | 10.4257 | 11.4049 | 10.6004 | 11.2302
Sm, Samarium (ppm) | 5.8589 | 5.47 | 6.2478 | 5.7037 | 6.014
Sn, Tin (ppm) | 11.6399 | 10.8761 | 12.4038 | 11.3294 | 11.9505
Sr, Strontium (ppm) | 52.4419 | 50.1481 | 54.7357 | 51.0039 | 53.8798
Ta, Tantalum (ppm) | 0.92955 | 0.85237 | 1.0067 | 0.88522 | 0.97387
Tb, Terbium (ppm) | 0.56329 | 0.47528 | 0.6513 | 0.52464 | 0.60194
Te, Tellurium (ppm) | 0.25997 | 0.21931 | 0.30064 | 0.2405 | 0.27945
Th, Thorium (ppm) | 15.0994 | 14.4754 | 15.7233 | 14.6816 | 15.5171
Ti, Titanium (wt.%) | 0.31503 | 0.29704 | 0.33302 | 0.30651 | 0.32354
Tl, Thallium (ppm) | 1.1576 | 1.109 | 1.2063 | 1.1267 | 1.1886
Tm, Thulium (ppm) | 0.23779 | 0.21247 | 0.2631 | 0.20571 | 0.26987
U, Uranium (ppm) | 2.4863 | 2.3543 | 2.6184 | 2.3871 | 2.5855
V, Vanadium (ppm) | 75.8243 | 74.1094 | 77.5392 | 74.0772 | 77.5714
W, Tungsten (ppm) | 42.5159 | 39.5552 | 45.4766 | 40.6148 | 44.4169
Y, Yttrium (ppm) | 14.3174 | 13.3947 | 15.2401 | 13.9535 | 14.6813
Yb, Ytterbium (ppm) | 1.6273 | 1.4243 | 1.8303 | 1.5379 | 1.7167
Zn, Zinc (ppm) | 44.0285 | 42.2223 | 45.8347 | 42.3344 | 45.7226
Zr, Zirconium (ppm) | 110.3413 | 103.9712 | 116.7113 | 105.1448 | 115.5377
Aqua Regia Digestion
Ag, Silver (ppm) | 0.27462 | 0.25229 | 0.29694 | 0.26054 | 0.2887
Al, Aluminium (wt.%) | 0.96715 | 0.90667 | 1.0276 | 0.93544 | 0.99886
As, Arsenic (ppm) | 231.2736 | 223.7124 | 238.8347 | 226.5375 | 236.0096
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 68.1944 | 61.3057 | 75.0831 | 66.109 | 70.2797
Be, Beryllium (ppm) | 0.68066 | 0.63024 | 0.73108 | 0.64671 | 0.7146
Bi, Bismuth (ppm) | 9.0964 | 8.5466 | 9.6462 | 8.7437 | 9.4491
Ca, Calcium (wt.%) | 0.16193 | 0.15416 | 0.1697 | 0.155 | 0.16886
Cd, Cadmium (ppm) | 0.034871 | 0.02462 | 0.045123 | IND | IND
Ce, Cerium (ppm) | 48.1594 | 44.2501 | 52.0687 | 46.8765 | 49.4424
Co, Cobalt (ppm) | 3.3635 | 3.0771 | 3.65 | 3.199 | 3.5281
Cr, Chromium (ppm) | 33.9228 | 32.3253 | 35.5204 | 32.6311 | 35.2146
Cs, Caesium (ppm) | 2.0892 | 1.9057 | 2.2727 | 2.0165 | 2.162
Cu, Copper (ppm) | 72.7162 | 69.6299 | 75.8026 | 71.0459 | 74.3866
Eu, Europium (ppm) | 0.59542 | 0.50577 | 0.68508 | 0.57124 | 0.61961
Fe, Iron (wt.%) | 2.2855 | 2.2128 | 2.3583 | 2.2393 | 2.3318
Ga, Gallium (ppm) | 2.8928 | 2.6681 | 3.1174 | 2.7474 | 3.0381
Hf, Hafnium (ppm) | 0.37185 | 0.33956 | 0.40414 | 0.34663 | 0.39706
Hg, Mercury (ppm) | 0.10045 | 0.071013 | 0.12989 | IND | IND
Ho, Holmium (ppm) | 0.2708 | 0.24125 | 0.30034 | 0.26084 | 0.28076
K, Potassium (wt.%) | 0.23826 | 0.22068 | 0.25584 | 0.22941 | 0.24712
La, Lanthanum (ppm) | 23.1701 | 21.0648 | 25.2754 | 22.5614 | 23.7788
Li, Lithium (ppm) | 8.0317 | 7.4491 | 8.6142 | 7.6731 | 8.3902
Mg, Magnesium (wt.%) | 0.1821 | 0.16938 | 0.19481 | 0.17277 | 0.19142
Mn, Manganese (wt.%) | 0.013385 | 0.012916 | 0.013854 | 0.013063 | 0.013707
Mo, Molybdenum (ppm) | 2.155 | 2.0287 | 2.2812 | 2.0365 | 2.2735
Na, Sodium (wt.%) | 0.019376 | 0.017667 | 0.021084 | 0.018377 | 0.020375
Nb, Niobium (ppm) | 0.12678 | 0.10528 | 0.14828 | IND | IND
Ni, Nickel (ppm) | 37.9302 | 36.425 | 39.4353 | 36.8142 | 39.0461
P, Phosphorus (wt.%) | 0.023103 | 0.021729 | 0.024476 | 0.022492 | 0.023713
Pb, Lead (ppm) | 12.9305 | 11.9922 | 13.8688 | 12.4373 | 13.4237
Pr, Praseodymium (ppm) | 6.4342 | 4.3623 | 8.506 | 6.2976 | 6.5708
Rb, Rubidium (ppm) | 20.9468 | 19.6395 | 22.2541 | 20.1433 | 21.7502
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.027092 | 0.022974 | 0.031209 | 0.025483 | 0.028701
Sb, Antimony (ppm) | 7.3509 | 6.6677 | 8.0341 | 7.0656 | 7.6362
Sc, Scandium (ppm) | 2.3265 | 2.1615 | 2.4915 | 2.2187 | 2.4344
Se, Selenium (ppm) | 0.45233 | 0.28977 | 0.6149 | IND | IND
Sn, Tin (ppm) | 1.6921 | 1.556 | 1.8282 | 1.5654 | 1.8188
Sr, Strontium (ppm) | 17.3283 | 15.2472 | 19.4093 | 16.6057 | 18.0508
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.37032 | 0.27508 | 0.46556 | 0.34794 | 0.3927
Te, Tellurium (ppm) | 0.22054 | 0.19543 | 0.24564 | 0.201 | 0.24007
Th, Thorium (ppm) | 10.7121 | 10.0809 | 11.3432 | 10.455 | 10.9691
Ti, Titanium (wt.%) | 0.014892 | 0.012774 | 0.01701 | 0.013941 | 0.015843
Tl, Thallium (ppm) | 0.42594 | 0.39546 | 0.45642 | 0.40811 | 0.44376
U, Uranium (ppm) | 1.3053 | 1.232 | 1.3786 | 1.2635 | 1.3471
V, Vanadium (ppm) | 22.0087 | 20.7945 | 23.2229 | 21.0089 | 23.0085
W, Tungsten (ppm) | 14.7639 | 13.2015 | 16.3264 | 14.0701 | 15.4578
Y, Yttrium (ppm) | 6.1998 | 5.788 | 6.6115 | 5.9932 | 6.4063
Zn, Zinc (ppm) | 29.9529 | 27.9491 | 31.9567 | 28.6948 | 31.2111
Zr, Zirconium (ppm) | 15.6975 | 14.7542 | 16.6408 | 15.0897 | 16.3052
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)
*Gross mass refers to the mass of the entire jar assembly, including jar base, jar lid and contents. These value ranges were developed using a ~40g empty jar mass but should be achievable for any jar-lid combination.

## Indicative Values
Table 2. Indicative Values for OREAS 263b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 19.1667 | Pt | ppb | < 5 | Se | ppm | 0.86026
Hg | ppm | 0.28583 | Re | ppm | 0.0014313
Aqua Regia Digestion
Dy | ppm | 1.6615 | In | ppm | 0.0235 | Pt | ppb | 2.1667
Er | ppm | 0.60794 | Lu | ppm | 0.064803 | Sm | ppm | 3.8198
Gd | ppm | 2.831 | Nd | ppm | 21.4051 | Tm | ppm | 0.075304
Ge | ppm | 0.0825 | Pd | ppb | < 10 | Yb | ppm | 0.52992
Borate Fusion XRF
Al2O3 | wt.% | 12.685 | Fe2O3 | wt.% | 4.0246 | S | wt.% | 0.027
As | ppm | 275 | K2O | wt.% | 2.52 | SiO2 | wt.% | 75.615
BaO | ppm | 541.5025 | MgO | wt.% | 0.955 | Sn | ppm | 65
CaO | wt.% | 0.41 | MnO | wt.% | 0.025 | Sr | ppm | 50
Cl | ppm | 40 | Na2O | wt.% | 0.175 | TiO2 | wt.% | 0.642
Co | ppm | 10 | Ni | ppm | 70 | V2O5 | ppm | 160.668
Cr2O3 | ppm | 146.15 | P2O5 | wt.% | 0.069894 | Zn | ppm | 45
Cu | ppm | 85 | Pb | ppm | 50 | Zr | ppm | 240
Thermogravimetry
LOI1000 | wt.% | 2.67
Infrared Combustion
C | wt.% | 0.06 | S | wt.% | 0.155
Laser Ablation ICP-MS
Ag | ppm | 0.3 | Hf | ppm | 6.98 | Sm | ppm | 6.535
As | ppm | 260 | Ho | ppm | 1.025 | Sn | ppm | 14.6
Ba | ppm | 494 | In | ppm | 0.05 | Sr | ppm | 54.9
Be | ppm | 2.4 | La | ppm | 42.4 | Ta | ppm | 1.14
Bi | ppm | 12 | Lu | ppm | 0.395 | Tb | ppm | 0.81
Cd | ppm | 0.075 | Mn | wt.% | 0.01785 | Te | ppm | < 0.2
Ce | ppm | 84.05 | Mo | ppm | 2.6 | Th | ppm | 15.6
Co | ppm | 5.3 | Nb | ppm | 13.25 | Ti | wt.% | 0.396
Cr | ppm | 98 | Nd | ppm | 34.8 | Tl | ppm | 0.4
Cs | ppm | 8.215 | Ni | ppm | 46 | Tm | ppm | 0.445
Cu | ppm | 74 | Pb | ppm | 20 | U | ppm | 2.875
Dy | ppm | 4.875 | Pr | ppm | 9.62 | V | ppm | 86.5
Er | ppm | 2.895 | Rb | ppm | 128.5 | W | ppm | 46.25
Eu | ppm | 1.07 | Re | ppm | < 0.01 | Y | ppm | 26.45
Ga | ppm | 16.3 | Sb | ppm | 13.35 | Yb | ppm | 2.865
Gd | ppm | 5.18 | Sc | ppm | 11.55 | Zn | ppm | 45
Ge | ppm | 1.6 | Se | ppm | < 5 | Zr | ppm | 243
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 263b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.21447 | 0.011864 | 0.19075 | 0.2382 | 0.17888 | 0.25007 | 5.532% | 11.06% | 16.6% | 0.20375 | 0.2252
PhotonAssay
Au, ppm | 0.21103 | 0.026044 | 0.15895 | 0.26312 | 0.1329 | 0.28917 | 12.34% | 24.68% | 37.02% | 0.20048 | 0.22159
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.19108 | 0.015914 | 0.15925 | 0.22291 | 0.14333 | 0.23882 | 8.329% | 16.66% | 24.99% | 0.18152 | 0.20063
Cyanide Leach
Au, ppm | 0.19098 | 0.015187 | 0.16061 | 0.22135 | 0.14542 | 0.23654 | 7.952% | 15.9% | 23.86% | 0.18143 | 0.20053
4-Acid Digestion
Ag, ppm | 0.30715 | 0.016693 | 0.27376 | 0.34053 | 0.25707 | 0.35723 | 5.435% | 10.87% | 16.3% | 0.29179 | 0.3225
Al, wt.% | 6.2983 | 0.21857 | 5.8612 | 6.7354 | 5.6426 | 6.954 | 3.47% | 6.941% | 10.41% | 5.9834 | 6.6132
As, ppm | 257.4637 | 7.9554 | 241.5528 | 273.3746 | 233.5974 | 281.33 | 3.09% | 6.18% | 9.27% | 244.5905 | 270.3369
Ba, ppm | 479.1494 | 13.9571 | 451.2352 | 507.0635 | 437.2781 | 521.0206 | 2.913% | 5.826% | 8.739% | 455.1919 | 503.1068
Be, ppm | 2.1588 | 0.13594 | 1.887 | 2.4307 | 1.751 | 2.5666 | 6.297% | 12.59% | 18.89% | 2.0509 | 2.2668
Bi, ppm | 10.1057 | 0.62395 | 8.8578 | 11.3536 | 8.2339 | 11.9776 | 6.174% | 12.35% | 18.52% | 9.6004 | 10.611
Ca, wt.% | 0.29448 | 0.0088934 | 0.2767 | 0.31227 | 0.2678 | 0.32116 | 3.02% | 6.04% | 9.06% | 0.27976 | 0.30921
Cd, ppm | 0.04369 | 0.017359 | 0.0089723 | 0.078407 | 0 | 0.095766 | 39.73% | 79.46% | 119.2% | 0.041505 | 0.045874
Ce, ppm | 82.2028 | 3.6938 | 74.8152 | 89.5904 | 71.1214 | 93.2842 | 4.494% | 8.987% | 13.48% | 78.0927 | 86.313
Co, ppm | 4.8092 | 0.2605 | 4.2881 | 5.3302 | 4.0276 | 5.5907 | 5.417% | 10.83% | 16.25% | 4.5687 | 5.0496
Cr, ppm | 78.1309 | 6.2793 | 65.5724 | 90.6895 | 59.2931 | 96.9688 | 8.037% | 16.07% | 24.11% | 74.2244 | 82.0375
Cs, ppm | 8.0141 | 0.3632 | 7.2877 | 8.7405 | 6.9245 | 9.1037 | 4.532% | 9.064% | 13.6% | 7.6134 | 8.4148
Cu, ppm | 74.9709 | 3.533 | 67.9048 | 82.037 | 64.3717 | 85.57 | 4.713% | 9.425% | 14.14% | 71.2223 | 78.7194
Dy, ppm | 3.2126 | 0.28179 | 2.649 | 3.7762 | 2.3672 | 4.058 | 8.771% | 17.54% | 26.31% | 3.052 | 3.3732
Er, ppm | 1.6664 | 0.2057 | 1.255 | 2.0778 | 1.0493 | 2.2835 | 12.34% | 24.69% | 37.03% | 1.5831 | 1.7497
Eu, ppm | 1.0244 | 0.081913 | 0.86053 | 1.1882 | 0.77862 | 1.2701 | 7.996% | 15.99% | 23.99% | 0.97314 | 1.0756
Fe, wt.% | 2.6821 | 0.098228 | 2.4857 | 2.8786 | 2.3874 | 2.9768 | 3.662% | 7.325% | 10.99% | 2.548 | 2.8162
Ga, ppm | 16.5685 | 0.60503 | 15.3585 | 17.7786 | 14.7534 | 18.3836 | 3.652% | 7.303% | 10.95% | 15.7401 | 17.397
Gd, ppm | 4.2644 | 0.2402 | 3.784 | 4.7448 | 3.5438 | 4.985 | 5.633% | 11.27% | 16.9% | 4.0512 | 4.4776
Ge, ppm | 0.11739 | 0.012075 | 0.093238 | 0.14154 | 0.081163 | 0.15361 | 10.29% | 20.57% | 30.86% | 0.11152 | 0.12326
Hf, ppm | 3.3742 | 0.20921 | 2.9558 | 3.7926 | 2.7466 | 4.0018 | 6.2% | 12.4% | 18.6% | 3.2055 | 3.5429
Ho, ppm | 0.57601 | 0.071774 | 0.43246 | 0.71956 | 0.36069 | 0.79133 | 12.46% | 24.92% | 37.38% | 0.54721 | 0.60481
In, ppm | 0.080845 | 0.0068705 | 0.067104 | 0.094586 | 0.060233 | 0.10146 | 8.498% | 17% | 25.5% | 0.076803 | 0.084887
K, wt.% | 2.0324 | 0.065133 | 1.9021 | 2.1626 | 1.837 | 2.2277 | 3.205% | 6.41% | 9.614% | 1.9307 | 2.134
La, ppm | 40.2916 | 2.7617 | 34.7682 | 45.815 | 32.0065 | 48.5767 | 6.854% | 13.71% | 20.56% | 38.277 | 42.3062
Li, ppm | 37.9046 | 1.7276 | 34.4494 | 41.3597 | 32.7218 | 43.0873 | 4.558% | 9.115% | 13.67% | 36.0093 | 39.7998
Lu, ppm | 0.2664 | 0.023454 | 0.21949 | 0.31331 | 0.19604 | 0.33676 | 8.804% | 17.61% | 26.41% | 0.25308 | 0.27972
Mg, wt.% | 0.52699 | 0.032298 | 0.46239 | 0.59159 | 0.4301 | 0.62388 | 6.129% | 12.26% | 18.39% | 0.50064 | 0.55334
Mn, wt.% | 0.017356 | 0.00070132 | 0.015954 | 0.018759 | 0.015252 | 0.01946 | 4.041% | 8.081% | 12.12% | 0.016489 | 0.018224
Mo, ppm | 2.3601 | 0.1731 | 2.0139 | 2.7063 | 1.8408 | 2.8794 | 7.334% | 14.67% | 22% | 2.2421 | 2.4781
Na, wt.% | 0.11487 | 0.0082139 | 0.098442 | 0.1313 | 0.090228 | 0.13951 | 7.151% | 14.3% | 21.45% | 0.10913 | 0.12061
Nb, ppm | 10.0686 | 1.2377 | 7.5933 | 12.544 | 6.3556 | 13.7816 | 12.29% | 24.58% | 36.88% | 9.5652 | 10.5721
Nd, ppm | 33.0779 | 1.7458 | 29.5863 | 36.5694 | 27.8405 | 38.3152 | 5.278% | 10.56% | 15.83% | 31.424 | 34.7318
Ni, ppm | 44.0176 | 1.2764 | 41.4647 | 46.5705 | 40.1883 | 47.847 | 2.9% | 5.8% | 8.7% | 41.8167 | 46.2185
P, wt.% | 0.029386 | 0.0010684 | 0.02725 | 0.031523 | 0.026181 | 0.032591 | 3.636% | 7.271% | 10.91% | 0.027917 | 0.030856
Pb, ppm | 19.3605 | 1.1255 | 17.1094 | 21.6115 | 15.9839 | 22.737 | 5.813% | 11.63% | 17.44% | 18.3925 | 20.3285
Pr, ppm | 9.184 | 0.61795 | 7.9481 | 10.4199 | 7.3302 | 11.0379 | 6.728% | 13.46% | 20.19% | 8.7248 | 9.6432
Rb, ppm | 131.5962 | 4.226 | 123.1441 | 140.0483 | 118.9181 | 144.2743 | 3.211% | 6.423% | 9.634% | 125.0164 | 138.176
S, wt.% | 0.028769 | 0.0050988 | 0.018571 | 0.038967 | 0.013473 | 0.044065 | 17.72% | 35.45% | 53.17% | 0.027331 | 0.030208
Sb, ppm | 11.3175 | 1.0656 | 9.1862 | 13.4487 | 8.1206 | 14.5143 | 9.416% | 18.83% | 28.25% | 10.7516 | 11.8833
Sc, ppm | 10.9153 | 0.56797 | 9.7793 | 12.0512 | 9.2114 | 12.6192 | 5.203% | 10.41% | 15.61% | 10.3695 | 11.461
Sm, ppm | 5.8589 | 0.37154 | 5.1158 | 6.602 | 4.7443 | 6.9735 | 6.341% | 12.68% | 19.02% | 5.5659 | 6.1518
Sn, ppm | 11.6399 | 1.0471 | 9.5457 | 13.7341 | 8.4986 | 14.7812 | 8.996% | 17.99% | 26.99% | 11.0579 | 12.2219
Sr, ppm | 52.4419 | 2.3526 | 47.7367 | 57.147 | 45.3841 | 59.4996 | 4.486% | 8.972% | 13.46% | 49.8198 | 55.064
Ta, ppm | 0.92955 | 0.044158 | 0.84123 | 1.0179 | 0.79707 | 1.062 | 4.75% | 9.501% | 14.25% | 0.88307 | 0.97603
Tb, ppm | 0.56329 | 0.08387 | 0.39555 | 0.73103 | 0.31168 | 0.8149 | 14.89% | 29.78% | 44.67% | 0.53513 | 0.59145
Te, ppm | 0.25997 | 0.033797 | 0.19238 | 0.32757 | 0.15858 | 0.36137 | 13% | 26% | 39% | 0.24698 | 0.27297
Th, ppm | 15.0994 | 0.70795 | 13.6835 | 16.5153 | 12.9755 | 17.2232 | 4.689% | 9.377% | 14.07% | 14.3444 | 15.8543
Ti, wt.% | 0.31503 | 0.02699 | 0.26105 | 0.36901 | 0.23406 | 0.396 | 8.568% | 17.14% | 25.7% | 0.29928 | 0.33078
Tl, ppm | 1.1576 | 0.072465 | 1.0127 | 1.3026 | 0.94025 | 1.375 | 6.26% | 12.52% | 18.78% | 1.0998 | 1.2155
Tm, ppm | 0.23779 | 0.015351 | 0.20709 | 0.26849 | 0.19174 | 0.28384 | 6.456% | 12.91% | 19.37% | 0.2259 | 0.24968
U, ppm | 2.4863 | 0.13385 | 2.2186 | 2.754 | 2.0848 | 2.8879 | 5.383% | 10.77% | 16.15% | 2.362 | 2.6106
V, ppm | 75.8243 | 1.8264 | 72.1715 | 79.4771 | 70.3451 | 81.3035 | 2.409% | 4.817% | 7.226% | 72.0331 | 79.6155
W, ppm | 42.5159 | 3.6499 | 35.2161 | 49.8156 | 31.5662 | 53.4655 | 8.585% | 17.17% | 25.75% | 40.3901 | 44.6416
Y, ppm | 14.3174 | 1.532 | 11.2534 | 17.3814 | 9.7213 | 18.9135 | 10.7% | 21.4% | 32.1% | 13.6015 | 15.0333
Yb, ppm | 1.6273 | 0.23744 | 1.1524 | 2.1022 | 0.91495 | 2.3396 | 14.59% | 29.18% | 43.77% | 1.5459 | 1.7086
Zn, ppm | 44.0285 | 2.099 | 39.8305 | 48.2265 | 37.7315 | 50.3255 | 4.767% | 9.535% | 14.3% | 41.8271 | 46.2299
Zr, ppm | 110.3413 | 9.0756 | 92.19 | 128.4925 | 83.1144 | 137.5681 | 8.225% | 16.45% | 24.68% | 104.8242 | 115.8583
Aqua Regia Digestion
Ag, ppm | 0.27462 | 0.020318 | 0.23398 | 0.31525 | 0.21366 | 0.33557 | 7.399% | 14.8% | 22.2% | 0.26089 | 0.28835
Al, wt.% | 0.96715 | 0.05997 | 0.84721 | 1.0871 | 0.78724 | 1.1471 | 6.201% | 12.4% | 18.6% | 0.91879 | 1.0155
As, ppm | 231.2736 | 7.4705 | 216.3326 | 246.2145 | 208.8621 | 253.685 | 3.23% | 6.46% | 9.69% | 219.7099 | 242.8372
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 68.1944 | 9.3965 | 49.4015 | 86.9873 | 40.005 | 96.3838 | 13.78% | 27.56% | 41.34% | 64.7847 | 71.6041
Be, ppm | 0.68066 | 0.046927 | 0.58681 | 0.77451 | 0.53988 | 0.82144 | 6.894% | 13.79% | 20.68% | 0.64663 | 0.71469
Bi, ppm | 9.0964 | 0.51398 | 8.0684 | 10.1244 | 7.5545 | 10.6384 | 5.65% | 11.3% | 16.95% | 8.6416 | 9.5512
Ca, wt.% | 0.16193 | 0.0051712 | 0.15159 | 0.17227 | 0.14642 | 0.17744 | 3.193% | 6.387% | 9.58% | 0.15383 | 0.17003
Cd, ppm | 0.034871 | 0.0070755 | 0.02072 | 0.049022 | 0.013645 | 0.056098 | 20.29% | 40.58% | 60.87% | 0.033128 | 0.036615
Ce, ppm | 48.1594 | 5.4646 | 37.2301 | 59.0887 | 31.7655 | 64.5533 | 11.35% | 22.69% | 34.04% | 45.7514 | 50.5674
Co, ppm | 3.3635 | 0.41133 | 2.5409 | 4.1862 | 2.1296 | 4.5975 | 12.23% | 24.46% | 36.69% | 3.1954 | 3.5317
Cr, ppm | 33.9228 | 1.4508 | 31.0212 | 36.8245 | 29.5703 | 38.2754 | 4.277% | 8.554% | 12.83% | 32.2267 | 35.619
Cs, ppm | 2.0892 | 0.22756 | 1.6341 | 2.5444 | 1.4065 | 2.7719 | 10.89% | 21.78% | 32.68% | 1.9848 | 2.1937
Cu, ppm | 72.7162 | 2.8686 | 66.979 | 78.4535 | 64.1103 | 81.3222 | 3.945% | 7.89% | 11.83% | 69.0804 | 76.3521
Eu, ppm | 0.59542 | 0.058825 | 0.47778 | 0.71307 | 0.41895 | 0.7719 | 9.879% | 19.76% | 29.64% | 0.56565 | 0.6252
Fe, wt.% | 2.2855 | 0.094447 | 2.0966 | 2.4744 | 2.0022 | 2.5689 | 4.132% | 8.265% | 12.4% | 2.1712 | 2.3998
Ga, ppm | 2.8928 | 0.25344 | 2.3859 | 3.3996 | 2.1324 | 3.6531 | 8.761% | 17.52% | 26.28% | 2.7481 | 3.0374
Hf, ppm | 0.37185 | 0.031212 | 0.30942 | 0.43427 | 0.27821 | 0.46548 | 8.394% | 16.79% | 25.18% | 0.35326 | 0.39044
Hg, ppm | 0.10045 | 0.025532 | 0.049385 | 0.15151 | 0.023852 | 0.17705 | 25.42% | 50.84% | 76.25% | 0.095427 | 0.10547
Ho, ppm | 0.2708 | 0.021521 | 0.22775 | 0.31384 | 0.20623 | 0.33536 | 7.947% | 15.89% | 23.84% | 0.25726 | 0.28434
K, wt.% | 0.23826 | 0.021503 | 0.19526 | 0.28127 | 0.17375 | 0.30277 | 9.025% | 18.05% | 27.07% | 0.22635 | 0.25018
La, ppm | 23.1701 | 3.1729 | 16.8242 | 29.516 | 13.6513 | 32.6889 | 13.69% | 27.39% | 41.08% | 22.0116 | 24.3286
Li, ppm | 8.0317 | 0.53265 | 6.9664 | 9.097 | 6.4337 | 9.6296 | 6.632% | 13.26% | 19.9% | 7.6301 | 8.4332
Mg, wt.% | 0.1821 | 0.014241 | 0.15362 | 0.21058 | 0.13938 | 0.22482 | 7.821% | 15.64% | 23.46% | 0.17299 | 0.1912
Mn, wt.% | 0.013385 | 0.00059963 | 0.012186 | 0.014584 | 0.011586 | 0.015184 | 4.48% | 8.96% | 13.44% | 0.012716 | 0.014054
Mo, ppm | 2.155 | 0.13217 | 1.8906 | 2.4193 | 1.7585 | 2.5515 | 6.133% | 12.27% | 18.4% | 2.0472 | 2.2627
Na, wt.% | 0.019376 | 0.0018647 | 0.015646 | 0.023105 | 0.013782 | 0.02497 | 9.624% | 19.25% | 28.87% | 0.018407 | 0.020345
Nb, ppm | 0.12678 | 0.021106 | 0.084565 | 0.16899 | 0.063459 | 0.1901 | 16.65% | 33.3% | 49.94% | 0.12044 | 0.13312
Ni, ppm | 37.9302 | 1.4282 | 35.0737 | 40.7866 | 33.6455 | 42.2148 | 3.765% | 7.531% | 11.3% | 36.0337 | 39.8267
P, wt.% | 0.023103 | 0.0019009 | 0.019301 | 0.026904 | 0.0174 | 0.028805 | 8.228% | 16.46% | 24.68% | 0.021948 | 0.024258
Pb, ppm | 12.9305 | 1.242 | 10.4464 | 15.4145 | 9.2044 | 16.6566 | 9.605% | 19.21% | 28.82% | 12.2839 | 13.577
Pr, ppm | 6.4342 | 1.5138 | 3.4066 | 9.4617 | 1.8928 | 10.9755 | 23.53% | 47.05% | 70.58% | 6.1125 | 6.7559
Rb, ppm | 20.9468 | 1.299 | 18.3488 | 23.5447 | 17.0499 | 24.8437 | 6.201% | 12.4% | 18.6% | 19.8994 | 21.9941
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.027092 | 0.0054247 | 0.016242 | 0.037941 | 0.010818 | 0.043366 | 20.02% | 40.05% | 60.07% | 0.025737 | 0.028446
Sb, ppm | 7.3509 | 1.266 | 4.819 | 9.8828 | 3.553 | 11.1488 | 17.22% | 34.44% | 51.67% | 6.9833 | 7.7184
Sc, ppm | 2.3265 | 0.17211 | 1.9823 | 2.6707 | 1.8102 | 2.8429 | 7.398% | 14.8% | 22.19% | 2.2102 | 2.4428
Se, ppm | 0.45233 | 0.088482 | 0.27537 | 0.6293 | 0.18689 | 0.71778 | 19.56% | 39.12% | 58.68% | 0.42972 | 0.47495
Sn, ppm | 1.6921 | 0.11696 | 1.4582 | 1.926 | 1.3412 | 2.043 | 6.912% | 13.82% | 20.74% | 1.6075 | 1.7767
Sr, ppm | 17.3283 | 3.2708 | 10.7867 | 23.8698 | 7.5159 | 27.1406 | 18.88% | 37.75% | 56.63% | 16.4618 | 18.1947
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.37032 | 0.071596 | 0.22713 | 0.51351 | 0.15553 | 0.58511 | 19.33% | 38.67% | 58% | 0.3518 | 0.38884
Te, ppm | 0.22054 | 0.020208 | 0.18012 | 0.26095 | 0.15991 | 0.28116 | 9.163% | 18.33% | 27.49% | 0.20951 | 0.23156
Th, ppm | 10.7121 | 0.94506 | 8.822 | 12.6022 | 7.8769 | 13.5473 | 8.822% | 17.64% | 26.47% | 10.1765 | 11.2477
Ti, wt.% | 0.014892 | 0.0036274 | 0.0076373 | 0.022147 | 0.0040099 | 0.025774 | 24.36% | 48.72% | 73.07% | 0.014147 | 0.015637
Tl, ppm | 0.42594 | 0.034311 | 0.35732 | 0.49456 | 0.32301 | 0.52887 | 8.055% | 16.11% | 24.17% | 0.40464 | 0.44724
U, ppm | 1.3053 | 0.094023 | 1.1173 | 1.4934 | 1.0232 | 1.5874 | 7.203% | 14.41% | 21.61% | 1.2401 | 1.3706
V, ppm | 22.0087 | 1.1131 | 19.7825 | 24.2349 | 18.6693 | 25.348 | 5.058% | 10.12% | 15.17% | 20.9082 | 23.1091
W, ppm | 14.7639 | 1.4113 | 11.9413 | 17.5866 | 10.53 | 18.9979 | 9.559% | 19.12% | 28.68% | 14.0258 | 15.5021
Y, ppm | 6.1998 | 0.3784 | 5.443 | 6.9566 | 5.0646 | 7.335 | 6.103% | 12.21% | 18.31% | 5.8898 | 6.5097
Zn, ppm | 29.9529 | 2.7389 | 24.4752 | 35.4307 | 21.7363 | 38.1696 | 9.144% | 18.29% | 27.43% | 28.4553 | 31.4506
Zr, ppm | 15.6975 | 1.0335 | 13.6305 | 17.7645 | 12.597 | 18.7979 | 6.584% | 13.17% | 19.75% | 14.9126 | 16.4824
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 263b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
NA | Unable to report due to QC failure (Lab 1.15)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 263b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST (SÖGÜT), Sögüt, Bilecik Province, Turkey
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Britannia Mining Solutions, Hamilton, Ontario, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Mineral Solutions (BVMS), Al Wadi District, Jeddah, Saudi Arabia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Limited, Manso Nkwanta, Ashanti Region, Ghana
Intertek Minerals Ltd, Bibiani, Western North Region, Ghana
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Laboratoire LABOMINE SARL, Agadir, Souss-Massa, Morocco
Labwest Minerals Analysis, Perth, WA, Australia
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Ghana Ltd, Obuasi, Ashanti, Ghana
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 263b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 16 | 1 | 267779 | 0.084472 | 0.20825 | 0.21665
2 | 1 | ANSLu | 1 | 1 | 3 | 267780 | 0.086747 | 0.24971 | 0.21886
2 | 1 | ANSLu | 1 | 6 | 6 | 267781 | 0.084756 | 0.20761 | 0.21662
2 | 1 | ANSLu | 1 | 7 | 10 | 267782 | 0.086439 | 0.20478 | 0.21647
2 | 1 | ANSLu | 1 | 20 | 12 | 267783 | 0.084744 | 0.21746 | 0.21714
2 | 1 | ANSLu | 1 | 8 | 7 | 267784 | 0.086505 | 0.22704 | 0.21765
2 | 1 | ANSLu | 1 | 4 | 3 | 267785 | 0.085268 | 0.20975 | 0.21673
2 | 1 | ANSLu | 1 | 15 | 4 | 267786 | 0.082862 | 0.21926 | 0.21724
2 | 1 | ANSLu | 1 | 10 | 10 | 267787 | 0.083226 | 0.21501 | 0.21701
2 | 1 | ANSLu | 1 | 13 | 9 | 267788 | 0.085597 | 0.20799 | 0.21664
2 | 1 | ANSLu | 1 | 17 | 11 | 267789 | 0.08216 | 0.19998 | 0.21621
2 | 1 | ANSLu | 1 | 5 | 2 | 267790 | 0.08445 | 0.20492 | 0.21647
2 | 1 | ANSLu | 1 | 18 | 9 | 267791 | 0.083838 | 0.21967 | 0.21726
2 | 1 | ANSLu | 1 | 14 | 4 | 267792 | 0.086649 | 0.20346 | 0.21639
2 | 1 | ANSLu | 1 | 19 | 1 | 267793 | 0.083831 | 0.24931 | 0.21883
2 | 1 | ANSLu | 1 | 12 | 12 | 267794 | 0.084746 | 0.21543 | 0.21703
2 | 1 | ANSLu | 1 | 11 | 5 | 267795 | 0.088049 | 0.2237 | 0.21747
2 | 1 | ANSLu | 1 | 3 | 8 | 267796 | 0.083413 | 0.21315 | 0.21691
2 | 1 | ANSLu | 1 | 2 | 7 | 267797 | 0.087121 | 0.22471 | 0.21753
2 | 1 | ANSLu | 1 | 9 | 6 | 267798 | 0.084716 | 0.22126 | 0.21734
Mean | [6 empty] | 0.084979 | 0.21712 | 0.21712
Median | [7 empty] | 0.21522 | 0.21702
SD | [7 empty] | 0.013449 | 0.00071581
RSD | [7 empty] | 6.194% | 0.3297%
RSD Upscale Factor | [7 empty] | 0.053223
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 263b (Certified Value 0.214 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.28 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 50g | 25g | 30g | 30g | 50g | 50g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 40g | 50g | 30g | 40g | 3
| 1 | 1 | 0.24971 | 0.19 | 0.206 | 0.20409 | 0.19 | 0.22464 | 0.23 | 0.223 | 0.21 | 0.22 | 0.23 | 0.2 | 0.21 | 0.23 | 0.22268 | 0.21 | 0.22 | 0.23 | 0.194 | 0.217 | 0.16 | 0.207 | 0.2 | 0.2 | 1
| 1 | 2 | 0.22471 | 0.22 | 0.2 | 0.20771 | 0.186 | 0.23136 | 0.22 | 0.216 | 0.21 | 0.24 | 0.24 | 0.22 | 0.2 | 0.23 | 0.22666 | 0.211 | 0.22 | 0.216 | 0.202 | 0.212 | 0.14 | 0.215 | 0.201 | 0.22 | #N/A
| 1 | 3 | 0.21315 | 0.2 | 0.219 | 0.20787 | 0.19 | 0.2072 | 0.23 | 0.249 | 0.2 | 0.22 | 0.22 | 0.21 | 0.21 | 0.24 | 0.21773 | 0.207 | 0.23 | 0.224 | 0.196 | 0.221 | 0.14 | 0.213 | 0.204 | 0.2 | 16
| 1 | 4 | 0.20975 | 0.2 | 0.209 | 0.2118 | 0.203 | 0.22656 | 0.22 | 0.221 | 0.21 | 0.23 | 0.21 | 0.27 | 0.21 | 0.23 | 0.22 | 0.207 | 0.22 | 0.214 | 0.207 | 0.22 | 0.15 | 0.213 | 0.204 | 0.19 | 0.21447
| 1 | 5 | 0.20492 | 0.19 | 0.221 | 0.20707 | 0.203 | 0.21216 | 0.22 | 0.222 | 0.2 | 0.23 | 0.22 | 0.21 | 0.21 | 0.23 | 0.22362 | 0.218 | 0.23 | 0.217 | 0.213 | 0.233 | 0.14 | 0.231 | 0.205 | 0.2 | 7
| 1 | 6 | 0.20761 | 0.22 | 0.218 | 0.21176 | 0.203 | 0.23136 | 0.22 | 0.21 | 0.21 | 0.23 | 0.22 | 0.22 | 0.21 | 0.23 | 0.21588 | 0.207 | 0.23 | 0.231 | 0.2 | 0.229 | 0.13 | 0.22 | 0.201 | 0.21
|  | 7 | 0.20478
|  | 8 | 0.22704
|  | 9 | 0.22126
|  | 10 | 0.21501
|  | 11 | 0.2237
|  | 12 | 0.21543
|  | 13 | 0.20799
|  | 14 | 0.20346
|  | 15 | 0.21926
|  | 16 | 0.20825
|  | 17 | 0.19998
|  | 18 | 0.21967
|  | 19 | 0.24931
|  | 20 | 0.21746
| Mean |  | 0.21712 | 0.20333 | 0.21217 | 0.20838 | 0.19583 | 0.22221 | 0.22333 | 0.2235 | 0.20667 | 0.22833 | 0.22333 | 0.22167 | 0.20833 | 0.23167 | 0.22109 | 0.21 | 0.225 | 0.222 | 0.202 | 0.222 | 0.14333 | 0.2165 | 0.2025 | 0.20333
| Median |  | 0.21522 | 0.2 | 0.2135 | 0.20779 | 0.1965 | 0.2256 | 0.22 | 0.2215 | 0.21 | 0.23 | 0.22 | 0.215 | 0.21 | 0.23 | 0.22134 | 0.2085 | 0.225 | 0.2205 | 0.201 | 0.2205 | 0.14 | 0.214 | 0.2025 | 0.2
| Std Dev. |  | 0.013449 | 0.013663 | 0.0084242 | 0.0029636 | 0.0079854 | 0.010184 | 0.005164 | 0.013398 | 0.005164 | 0.0075277 | 0.010328 | 0.024833 | 0.0040825 | 0.0040825 | 0.0039866 | 0.0042895 | 0.0054772 | 0.0074027 | 0.0070711 | 0.007746 | 0.010328 | 0.0082401 | 0.0020736 | 0.010328
| Rel.Std.Dev. |  | 6.194% | 6.719% | 3.971% | 1.422% | 4.078% | 4.583% | 2.312% | 5.995% | 2.499% | 3.297% | 4.624% | 11.2% | 1.96% | 1.762% | 1.803% | 2.043% | 2.434% | 3.335% | 3.501% | 3.489% | 7.206% | 3.806% | 1.024% | 5.079%
| PDM3 |  | 1.235% | -5.194% | -1.076% | -2.84% | -8.691% | 3.609% | 4.131% | 4.209% | -3.64% | 6.462% | 4.131% | 3.354% | -2.863% | 8.016% | 3.087% | -2.086% | 4.908% | 3.509% | -5.816% | 3.509% | -33.17% | 0.9447% | -5.583% | -5.194%
| Z-Score (Absolute) |  | NA | 1.09 | 0.36 | 0.67 | 1.72 | 0.47 | 0.57 | 0.58 | 0.82 | 0.98 | 0.57 | 0.43 | 0.68 | 1.26 | 0.38 | 0.54 | 0.71 | 0.46 | 1.2 | 0.46 | 6.08 | 0 | 1.16 | 1.09
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 263b (Certified Value 0.211 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | 6.13 | 6.14 | 6.15 | 6.16 | 6.17 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | [17 empty] | 3
| 1 | 1 | 0.24971 | 0.2 | 0.2 | 0.237 | 0.21 | 0.16 | 0.208 | 0.236 | 0.255 | 0.206 | 0.21 | 0.221 | 0.194 | 0.213 | 0.181 | 0.221 | 0.2 | 0.225 | 1
| 1 | 2 | 0.22471 | 0.22 | 0.18 | 0.198 | 0.24 | 0.22 | 0.192 | 0.178 | 0.255 | 0.21 | 0.17 | 0.194 | 0.213 | 0.216 | 0.19 | 0.221 | 0.3 | 0.249 | #N/A
| 1 | 3 | 0.21315 | [16 empty] | 0.229 | 16
| 1 | 4 | 0.20975 | [16 empty] | 0.188 | 0.21103
| 1 | 5 | 0.20492 | 0.23 | 0.17 | 0.239 | 0.17 | 0.13 | 0.222 | 0.208 | 0.226 | 0.212 | 0.21 | 0.208 | 0.197 | 0.221 | 0.212 | 0.164 | 0.21 | 0.183 | 9
| 1 | 6 | 0.20761 | 0.28 | 0.23 | 0.222 | 0.18 | 0.26 | 0.196 | 0.194 | 0.203 | 0.216 | 0.25 | 0.268 | 0.191 | 0.196 | 0.24 | 0.164 | 0.27 | 0.216
| 1 | 7 | 0.20478 | [16 empty] | 0.227
| 1 | 8 | 0.22704 | [16 empty] | 0.181
| 1 | 9 | 0.22126 | 0.19 | 0.22 | 0.223 | 0.19 | 0.19 | 0.204 | 0.225 | 0.191 | 0.209 | 0.23 | 0.216 | 0.215 | 0.197 | 0.221 | 0.181 | 0.25 | 0.232
| 1 | 10 | 0.21501 | 0.24 | 0.19 | 0.233 | 0.2 | 0.22 | 0.181 | 0.222 | 0.236 | 0.235 | 0.23 | 0.171 | 0.213 | 0.22 | 0.205 | 0.181 | 0.2 | 0.234
| 1 | 11 | 0.2237 | [16 empty] | 0.249
| 1 | 12 | 0.21543 | [16 empty] | 0.252
|  | 13 | 0.20799
|  | 14 | 0.20346
|  | 15 | 0.21926
|  | 16 | 0.20825
|  | 17 | 0.19998
|  | 18 | 0.21967
|  | 19 | 0.24931
|  | 20 | 0.21746
| Mean |  | 0.21712 | 0.22667 | 0.19833 | 0.22533 | 0.19833 | 0.19667 | 0.2005 | 0.2105 | 0.22767 | 0.21467 | 0.21667 | 0.213 | 0.20383 | 0.2105 | 0.20817 | 0.18867 | 0.23833 | 0.22208
| Median |  | 0.21522 | 0.225 | 0.195 | 0.228 | 0.195 | 0.205 | 0.2 | 0.215 | 0.231 | 0.211 | 0.22 | 0.212 | 0.205 | 0.2145 | 0.2085 | 0.181 | 0.23 | 0.228
| Std Dev. |  | 0.013449 | 0.032042 | 0.023166 | 0.015135 | 0.024833 | 0.046762 | 0.014167 | 0.021575 | 0.02653 | 0.010501 | 0.027325 | 0.032398 | 0.010962 | 0.011221 | 0.021311 | 0.026174 | 0.041673 | 0.025368
| Rel.Std.Dev. |  | 6.194% | 14.14% | 11.68% | 6.717% | 12.52% | 23.78% | 7.066% | 10.25% | 11.65% | 4.892% | 12.61% | 15.21% | 5.378% | 5.33% | 10.24% | 13.87% | 17.49% | 11.42%
| PDM3 |  | 2.885% | 7.407% | -6.018% | 6.776% | -6.018% | -6.808% | -4.992% | -0.2532% | 7.881% | 1.721% | 2.669% | 0.9315% | -3.412% | -0.2532% | -1.359% | -10.6% | 12.94% | 5.236%
| Z-Score (Absolute) |  | NA | 0.94 | 0.71 | 0.86 | 0.71 | 0.81 | 0.58 | 0 | 1 | 0.24 | 0.36 | 0.15 | 0.39 | 0 | 0.14 | 1.27 | 1.62 | 0.67
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 263b (Certified Value 0.191 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*OES | AR*AAS | AR*AAS | AR*MS | AR*AAS | 3
|  |  | 0.085g | 50g | 25g | 30g | 50g | 30g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 10g | 30g | 15g | 10g | 15g | 25g | 30g | 10g | 40g | 25g | 3
| 1 | 1 | 0.24971 | 0.21 | 0.19 | 0.22837 | 0.18 | 0.16 | 0.19 | 0.203 | 0.19 | 0.184 | 0.179 | 0.177 | < 0.5 | 0.49 | 0.2 | 0.2137 | 0.195 | 0.1941 | 0.17 | 0.22 | 0.2 | 0.183 | 0.2 | 1
| 1 | 2 | 0.22471 | 0.21 | 0.19 | 0.23071 | 0.2 | 0.166 | 0.2 | 0.193 | 0.21 | 0.182 | 0.175 | 0.18 | < 0.5 | 0.209 | 0.2 | 0.2121 | 0.187 | 0.1964 | 0.19 | 0.16 | 0.19 | 0.184 | 0.22 | #N/A
| 1 | 3 | 0.21315 | 0.2 | 0.19 | 0.23075 | 0.17 | 0.16 | 0.2 | 0.175 | 0.18 | 0.178 | 0.179 | 0.175 | < 0.5 | 0.207 | 0.19 | 0.1988 | 0.18 | 0.1901 | 0.188 | 0.17 | 0.19 | 0.187 | 0.23 | 16
| 1 | 4 | 0.20975 | 0.19 | 0.19 | 0.21085 | 0.165 | 0.167 | 0.2 | 0.204 | 0.19 | 0.187 | 0.177 | 0.175 | < 0.5 | 0.197 | 0.19 | 0.2323 | 0.18 | 0.2077 | 0.194 | 0.18 | 0.22 | 0.183 | 0.2 | 0.19108
| 1 | 5 | 0.20492 | 0.19 | 0.18 | 0.21935 | 0.2 | 0.155 | 0.19 | 0.194 | 0.19 | 0.188 | 0.171 | 0.175 | < 0.5 | 0.197 | 0.2 | 0.2103 | 0.18 | 0.1907 | 0.215 | 0.16 | 0.2 | 0.183 | 0.2 | 11
| 1 | 6 | 0.20761 | 0.2 | 0.19 | 0.20199 | 0.17 | 0.165 | 0.2 | 0.197 | 0.19 | 0.18 | 0.176 | 0.174 | < 0.5 | 0.201 | 0.19 | 0.1919 | 0.176 | 0.2004 | 0.176 | 0.18 | 0.21 | 0.181 | 0.2
|  | 7 | 0.20478
|  | 8 | 0.22704
|  | 9 | 0.22126
|  | 10 | 0.21501
|  | 11 | 0.2237
|  | 12 | 0.21543
|  | 13 | 0.20799
|  | 14 | 0.20346
|  | 15 | 0.21926
|  | 16 | 0.20825
|  | 17 | 0.19998
|  | 18 | 0.21967
|  | 19 | 0.24931
|  | 20 | 0.21746
| Mean |  | 0.21712 | 0.2 | 0.18833 | 0.22034 | 0.18083 | 0.16217 | 0.19667 | 0.19433 | 0.19167 | 0.18317 | 0.17617 | 0.176 |  | 0.25017 | 0.195 | 0.20985 | 0.183 | 0.19657 | 0.18883 | 0.17833 | 0.20167 | 0.1835 | 0.20833
| Median |  | 0.21522 | 0.2 | 0.19 | 0.22386 | 0.175 | 0.1625 | 0.2 | 0.1955 | 0.19 | 0.183 | 0.1765 | 0.175 |  | 0.204 | 0.195 | 0.2112 | 0.18 | 0.19525 | 0.189 | 0.175 | 0.2 | 0.183 | 0.2
| Std Dev. |  | 0.013449 | 0.0089443 | 0.0040825 | 0.0119 | 0.015626 | 0.0046224 | 0.005164 | 0.010501 | 0.0098319 | 0.00392 | 0.0029944 | 0.0021909 |  | 0.1176 | 0.0054772 | 0.013917 | 0.0068702 | 0.0066482 | 0.015702 | 0.022286 | 0.01169 | 0.0019748 | 0.013292
| Rel.Std.Dev. |  | 6.194% | 4.472% | 2.168% | 5.401% | 8.641% | 2.85% | 2.626% | 5.403% | 5.13% | 2.14% | 1.7% | 1.245% |  | 47.01% | 2.809% | 6.632% | 3.754% | 3.382% | 8.316% | 12.5% | 5.797% | 1.076% | 6.38%
| PDM3 |  | 13.63% | 4.67% | -1.436% | 15.31% | -5.361% | -15.13% | 2.925% | 1.704% | 0.3085% | -4.14% | -7.803% | -7.891% |  | 30.92% | 2.053% | 9.825% | -4.227% | 2.873% | -1.174% | -6.669% | 5.542% | -3.965% | 9.031%
| Z-Score (Absolute) |  | NA | 0.48 | 0.32 | 1.86 | 0.83 | 2.1 | 0.25 | 0.09 | 0.09 | 0.67 | 1.14 | 1.16 | 3.88 | 3.89 | 0.14 | 1.15 | 0.68 | 0.24 | 0.28 | 1 | 0.59 | 0.65 | 1.04
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 263b (Certified Value 0.191 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.22 | 1.23 | 1.27 | 1.29 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 20g | 15g | 05g | 30g | 200g | 50g | 200g | 50g | 3
| 1 | 1 | 0.24971 | 0.2 | 0.18 | 0.15011 | 0.193 | 0.21308 | 0.19 | 0.192 | 0.2 | 0.19 | 0.11 | 0.19 | 0.21 | 0.18896 | 0.17 | 0.2 | 0.17 | 0.182 | 0.195 | 0.18 | 0.17 | 1
| 1 | 2 | 0.22471 | 0.2 | 0.17 | 0.15578 | 0.19 | 0.2017 | 0.23 | 0.182 | 0.19 | 0.19 | 0.15 | 0.23 | 0.19 | 0.17035 | 0.18 | 0.2 | 0.17 | 0.187 | 0.19 | 0.18 | 0.18 | #N/A
| 1 | 3 | 0.21315 | 0.2 | 0.18 | 0.15473 | 0.203 | 0.20502 | 0.23 | 0.188 | 0.2 | 0.22 | 0.12 | 0.13 | 0.19 | 0.1796 | 0.18 | 0.2 | 0.17 | 0.187 | 0.195 | 0.18 | 0.19 | 16
| 1 | 4 | 0.20975 | 0.2 | 0.17 | 0.15794 | 0.209 | 0.19674 | 0.2 | 0.186 | 0.2 | 0.2 | 0.12 | 0.15 | 0.19 | 0.19272 | 0.17 | 0.2 | 0.16 | 0.193 | 0.199 | 0.18 | 0.19 | 0.19098
| 1 | 5 | 0.20492 | 0.2 | 0.18 | 0.15836 | 0.189 | 0.2192 | 0.2 | 0.185 | 0.19 | 0.2 | 0.17 | 0.17 | 0.2 | 0.19161 | 0.17 | 0.2 | 0.17 | 0.186 | 0.201 | 0.18 | 0.23 | 13
| 1 | 6 | 0.20761 | 0.2 | 0.18 | 0.15597 | 0.208 | 0.20856 | 0.21 | 0.192 | 0.2 | 0.19 | 0.13 | 0.09 | 0.21 | 0.16438 | 0.19 | 0.2 | 0.17 | 0.192 | 0.2 | 0.17 | 0.18
|  | 7 | 0.20478
|  | 8 | 0.22704
|  | 9 | 0.22126
|  | 10 | 0.21501
|  | 11 | 0.2237
|  | 12 | 0.21543
|  | 13 | 0.20799
|  | 14 | 0.20346
|  | 15 | 0.21926
|  | 16 | 0.20825
|  | 17 | 0.19998
|  | 18 | 0.21967
|  | 19 | 0.24931
|  | 20 | 0.21746
| Mean |  | 0.21712 | 0.2 | 0.17667 | 0.15548 | 0.19867 | 0.20738 | 0.21 | 0.1875 | 0.19667 | 0.19833 | 0.13333 | 0.16 | 0.19833 | 0.18127 | 0.17667 | 0.2 | 0.16833 | 0.18783 | 0.19667 | 0.17833 | 0.19
| Median |  | 0.21522 | 0.2 | 0.18 | 0.15588 | 0.198 | 0.20679 | 0.205 | 0.187 | 0.2 | 0.195 | 0.125 | 0.16 | 0.195 | 0.18428 | 0.175 | 0.2 | 0.17 | 0.187 | 0.197 | 0.18 | 0.185
| Std Dev. |  | 0.013449 | 3.0405e-17 | 0.005164 | 0.0029708 | 0.0090921 | 0.008058 | 0.016733 | 0.0039875 | 0.005164 | 0.01169 | 0.022509 | 0.04858 | 0.0098319 | 0.011869 | 0.008165 | 3.0405e-17 | 0.0040825 | 0.0040702 | 0.0041312 | 0.0040825 | 0.020976
| Rel.Std.Dev. |  | 6.194% | 1.52e-14% | 2.923% | 1.911% | 4.577% | 3.886% | 7.968% | 2.127% | 2.626% | 5.894% | 16.88% | 30.36% | 4.957% | 6.548% | 4.622% | 1.52e-14% | 2.425% | 2.167% | 2.101% | 2.289% | 11.04%
| PDM3 |  | 13.69% | 4.723% | -7.495% | -18.59% | 4.024% | 8.588% | 9.959% | -1.823% | 2.977% | 3.85% | -30.18% | -16.22% | 3.85% | -5.084% | -7.495% | 4.723% | -11.86% | -1.648% | 2.977% | -6.622% | -0.5135%
| Z-Score (Absolute) |  | NA | 0.69 | 0.76 | 2.08 | 0.61 | 1.15 | 1.31 | 0.09 | 0.48 | 0.59 | 3.46 | 1.8 | 0.59 | 0.48 | 0.76 | 0.69 | 1.28 | 0.07 | 0.48 | 0.66 | 0.07
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 263b (Certified Value 0.307 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.31 | 0.2867 | < 0.3 | < 0.5 | 0.29 | 0.32 | 0.31 | 0.3 | 0.26 | 0.32884 | < 0.3 | 0.2 | 0.293 | 0.32 | 0.296 | 0.3 | < 0.5 | 0.4 | 0.459 |  | 1
| 1 | 2 | 0.3 | 0.3028 | < 0.3 | < 0.5 | 0.3 | 0.32 | 0.31 | 0.31 | 0.29 | 0.30252 | < 0.3 | 0.3 | 0.312 | 0.3 | 0.316 | 0.3 | < 0.5 | 0.2 | 0.509 |  | 22
| 1 | 3 | 0.34 | 0.29972 | 0.5 | < 0.5 | 0.3 | 0.33 | 0.33 | 0.29 | 0.29 | 0.31096 | < 0.3 | 0.3 | 0.314 | 0.3 | 0.296 | 0.29 | < 0.5 | 0.2 | 0.592 |  | 16
| 1 | 4 | 0.3 | 0.29454 | 0.5 | < 0.5 | 0.3 | 0.32 | 0.32 | 0.27 | 0.33 | 0.31635 | < 0.3 | 0.3 | 0.367 | 0.33 | 0.325 | 0.31 | < 0.5 | 0.2 | 0.725 |  | 0.30715
| 1 | 5 | 0.31 | 0.29868 | 0.6 | < 0.5 | 0.31 | 0.29 | 0.28 | 0.29 | 0.28 | 0.31469 | < 0.3 | 0.3 | 0.361 | 0.3 | 0.298 | 0.33 | < 0.5 | 0.2 | 1.217 |  | 15
| 1 | 6 | 0.31 | 0.30585 | < 0.3 | < 0.5 | 0.32 | 0.32 | 0.47 | 0.31 | 0.29 | 0.33272 | 0.3 | 0.3 | 0.301 | 0.32 | 0.312 | 0.32 | < 0.5 | 0.2 | 0.657
| Mean |  | 0.31167 | 0.29805 | 0.53333 |  | 0.30333 | 0.31667 | 0.33667 | 0.295 | 0.29 | 0.31768 | 0.3 | 0.28333 | 0.32467 | 0.31167 | 0.30717 | 0.30833 |  | 0.23333 | 0.69317
| Median |  | 0.31 | 0.2992 | 0.5 |  | 0.3 | 0.32 | 0.315 | 0.295 | 0.29 | 0.31552 | 0.3 | 0.3 | 0.313 | 0.31 | 0.305 | 0.305 |  | 0.2 | 0.6245
| Std Dev. |  | 0.01472 | 0.006748 | 0.057735 |  | 0.010328 | 0.013663 | 0.067429 | 0.015166 | 0.022804 | 0.011282 |  | 0.040825 | 0.031462 | 0.013292 | 0.012271 | 0.01472 |  | 0.08165 | 0.27412
| Rel.Std.Dev. |  | 4.723% | 2.264% | 10.83% |  | 3.405% | 4.315% | 20.03% | 5.141% | 7.863% | 3.551% |  | 14.41% | 9.691% | 4.265% | 3.995% | 4.774% |  | 34.99% | 39.55%
| PDM3 |  | 1.472% | -2.962% | 73.64% |  | -1.242% | 3.099% | 9.611% | -3.955% | -5.583% | 3.429% | -2.327% | -7.753% | 5.704% | 1.472% | 0.006449% | 0.3863% |  | -24.03% | 125.7%
| Z-Score (Absolute) |  | 0.11 | 0.82 | NA | NA | 0.46 | 0.46 | 1.83 | 1.03 | 1.37 | 0.53 | NA | NA | 1.01 | 0.11 | 0.19 | 0.11 | 4.12 | NA | 26.33
Results from laboratories 1.07, 1.08, 1.15, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 263b (Certified Value 6.3 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 6.2343 | 6.0509 | 6.57 | 6.48 | 6.41 | 6.06 | 6.03 | 6.18 | 6.26 | 6.2 | 6.4796 | 5.8839 | 6.28 | 6.8 | 6.4083 | 6.16 | 6.51 | 6.37 | 6.2 | 6.5649 | 1
| 1 | 2 | 6.3112 | 6.1148 | 6.5626 | 6.56 | 6.36 | 6.03 | 6.17 | 6.18 | 6.21 | 6.36 | 6.399 | 5.9042 | 6.09 | 6.96 | 6.5408 | 5.94 | 6.35 | 6.28 | 6.21 | 6.5741 | #N/A
| 1 | 3 | 6.2714 | 6.0882 | 6.5707 | 6.62 | 6.45 | 6.08 | 6.07 | 6.22 | 6.21 | 6.18 | 6.393 | 5.7722 | 6.11 | 6.14 | 6.4085 | 6.03 | 6 | 6.33 | 6.28 | 6.4552 | 16
| 1 | 4 | 6.2557 | 6.2678 | 6.5781 | 6.62 | 6.39 | 6.14 | 6.11 | 6.29 | 6.29 | 6.41 | 6.4741 | 5.8555 | 6.03 | 7.97 | 6.5582 | 5.99 | 6.36 | 6.3 | 6.23 | 6.5508 | 6.2983
| 1 | 5 | 6.3471 | 6.1767 | 6.564 | 6.63 | 6.39 | 6.02 | 6.19 | 6.17 | 6.32 | 6.32 | 6.4992 | 5.9459 | 6.2 | 6.51 | 6.4782 | 6.16 | 6.77 | 6.26 | 6.21 | 6.7093 | 16
| 1 | 6 | 6.3432 | 6.1079 | 6.57 | 6.57 | 6.43 | 6.14 | 6.11 | 6.17 | 6.35 | 6.32 | 6.4574 | 5.8794 | 6.06 | 6.1 | 6.5275 | 6.3 | 6.22 | 6.26 | 6.24 | 6.6498
| Mean |  | 6.2938 | 6.1344 | 6.5692 | 6.58 | 6.405 | 6.0783 | 6.1133 | 6.2017 | 6.2733 | 6.2983 | 6.4504 | 5.8735 | 6.1283 | 6.7467 | 6.4869 | 6.0967 | 6.3683 | 6.3 | 6.2283 | 6.584
| Median |  | 6.2913 | 6.1113 | 6.57 | 6.595 | 6.4 | 6.07 | 6.11 | 6.18 | 6.275 | 6.32 | 6.4657 | 5.8817 | 6.1 | 6.655 | 6.5029 | 6.095 | 6.355 | 6.29 | 6.22 | 6.5695
| Std Dev. |  | 0.047069 | 0.077152 | 0.0055203 | 0.056921 | 0.032094 | 0.052313 | 0.059889 | 0.047081 | 0.057504 | 0.090425 | 0.044236 | 0.05816 | 0.09411 | 0.69096 | 0.066389 | 0.13397 | 0.26057 | 0.043359 | 0.029269 | 0.087381
| Rel.Std.Dev. |  | 0.7479% | 1.258% | 0.08403% | 0.8651% | 0.5011% | 0.8606% | 0.9796% | 0.7592% | 0.9166% | 1.436% | 0.6858% | 0.9902% | 1.536% | 10.24% | 1.023% | 2.197% | 4.092% | 0.6882% | 0.4699% | 1.327%
| PDM3 |  | -0.07113% | -2.603% | 4.302% | 4.473% | 1.694% | -3.492% | -2.937% | -1.534% | -0.3964% | 0.0005824% | 2.415% | -6.744% | -2.699% | 7.119% | 2.995% | -3.201% | 1.112% | 0.02704% | -1.111% | 4.537%
| Z-Score (Absolute) |  | 0.01 | 0.66 | 1.12 | 1.16 | 0.45 | 0.89 | 0.75 | 0.39 | 0.09 | 0.01 | 0.63 | 1.73 | 0.69 | 1.84 | 0.78 | 0.82 | 0.3 | 0.02 | 0.28 | 1.18
| Analytical results for As in OREAS 263b (Certified Value 257 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 0
0
| 1 | 1 | 261.9 | 253.0251 | 269 | 261 | 247 | 253 | 262 | 251 | 247 | 247.9383 | 228.9 | 254 | 251.7 | 260.9 | 268.6 | 237 | 196.4 | 260 | 318.845 |  | 1
| 1 | 2 | 259.3 | 251.4851 | 269 | 260 | 246 | 261 | 264 | 254 | 253 | 251.4753 | 227.7 | 254 | 254.7 | 263.3 | 266.9 | 248 | 196.5 | 264 | 317.696 |  | 10
| 1 | 3 | 269.8 | 250.3309 | 264 | 261 | 243 | 256 | 262 | 254 | 248 | 247.975 | 223.8 | 261 | 258.1 | 257.5 | 267.7 | 241 | 198.8 | 264 | 308.934 |  | 16
| 1 | 4 | 265.5 | 253.8874 | 266 | 259 | 248 | 259 | 266 | 256 | 256 | 252.203 | 224.8 | 257 | 264.8 | 258.3 | 267.8 | 242 | 198.8 | 261 | 287.209 |  | 257.4637
| 1 | 5 | 259.8 | 255.2784 | 272 | 256 | 246 | 260 | 264 | 254 | 248 | 252.1644 | 226.3 | 250 | 271.5 | 260.4 | 274 | 257 | 196.5 | 262 | 324.467 |  | 17
| 1 | 6 | 260.6 | 250.0147 | 269 | 259 | 251 | 256 | 264 | 256 | 254 | 250.3361 | 226.3 | 256 | 269 | 261 | 272.3 | 238 | 193.3 | 265 | 324.307
| Mean |  | 262.8167 | 252.3369 | 268.1667 | 259.3333 | 246.8333 | 257.5 | 263.6667 | 254.1667 | 251 | 250.3487 | 226.3 | 255.3333 | 261.6333 | 260.2333 | 269.55 | 243.8333 | 196.7167 | 262.6667 | 313.5763
| Median |  | 261.25 | 252.2551 | 269 | 259.5 | 246.5 | 257.5 | 264 | 254 | 250.5 | 250.9057 | 226.3 | 255 | 261.45 | 260.65 | 268.2 | 241.5 | 196.5 | 263 | 318.2705
| Std Dev. |  | 4.0789 | 2.0819 | 2.7869 | 1.8619 | 2.6394 | 3.0166 | 1.5055 | 1.8348 | 3.7947 | 1.9723 | 1.8558 | 3.6697 | 8.0109 | 2.082 | 2.8905 | 7.5211 | 2.0272 | 1.9664 | 14.1087
| Rel.Std.Dev. |  | 1.552% | 0.8251% | 1.039% | 0.718% | 1.069% | 1.172% | 0.571% | 0.7219% | 1.512% | 0.7878% | 0.8201% | 1.437% | 3.062% | 0.8% | 1.072% | 3.085% | 1.031% | 0.7486% | 4.499%
| PDM3 |  | 2.079% | -1.991% | 4.157% | 0.7262% | -4.129% | 0.01411% | 2.409% | -1.281% | -2.511% | -2.764% | -12.1% | -0.8274% | 1.62% | 1.076% | 4.694% | -5.294% | -23.59% | 2.021% | 21.79%
| Z-Score (Absolute) |  | 0.58 | 0.56 | 1.17 | 0.2 | 1.17 | 0 | 0.67 | 0.36 | 0.71 | 0.78 | 3.41 | 0.24 | 0.45 | 0.3 | 1.32 | 1.49 | 6.65 | 0.56 | 6.13
| Analytical results for B in OREAS 263b (Indicative Value 19.2 ppm) | [21 empty] | Indicative
B | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.08 | [19 empty] | ppm
|  |  | 4A*OES | [19 empty] | 1
1
| 1 | 1 | 19 | [19 empty] | 1
| 1 | 2 | 21 | [19 empty] | 1
| 1 | 3 | 18 | [19 empty] | 16
| 1 | 4 | 20 | [19 empty] | 19.1667
| 1 | 5 | 19 | [19 empty] | 7
| 1 | 6 | 18
| Mean |  | 19.1667
| Median |  | 19
| Std Dev. |  | 1.169
| Rel.Std.Dev. |  | 6.099%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 263b (Certified Value 479 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 483.1 | 476.0229 | 456.7371 | 498 | 504 | 460 | 480 | 470 | 480 | 480 | 477.5877 | 453 | 472 | 490 | 521.1 | 484 | 528 | 497.2 | 473 | 473.918 | 1
| 1 | 2 | 476.7 | 475.9865 | 461.413 | 504 | 501 | 460 | 490 | 470 | 470 | 480 | 464.9554 | 446 | 472 | 493 | 515.9 | 476 | 552 | 487.4 | 479 | 477.713 | 24
| 1 | 3 | 495.7 | 475.9865 | 461.5031 | 490 | 500 | 470 | 480 | 470 | 470 | 470 | 468.9146 | 450 | 473 | 487 | 517.3 | 475 | 531 | 490.5 | 485 | 474.712 | 16
| 1 | 4 | 490 | 478.6835 | 461.7019 | 497 | 500 | 470 | 490 | 470 | 480 | 500 | 483.1263 | 447 | 462 | 520 | 515.9 | 478 | 548 | 481.9 | 485 | 476.147 | 479.1494
| 1 | 5 | 494.8 | 483.3923 | 458.6997 | 485 | 499 | 460 | 490 | 470 | 480 | 480 | 481.649 | 450 | 470 | 506 | 513.3 | 492 | 540 | 482.6 | 475 | 481.672 | 18
| 1 | 6 | 490.8 | 473.0989 | 463.981 | 501 | 498 | 470 | 490 | 470 | 480 | 490 | 487.3784 | 456 | 472 | 497 | 521.9 | 478 | 542 | 485.5 | 474 | 479.953
| Mean |  | 488.5167 | 477.1951 | 460.6726 | 495.8333 | 500.3333 | 465 | 486.6667 | 470 | 476.6667 | 483.3333 | 477.2686 | 450.3333 | 470.1667 | 498.8333 | 517.5667 | 480.5 | 540.1667 | 487.5167 | 478.5 | 477.3525
| Median |  | 490.4 | 476.0047 | 461.4581 | 497.5 | 500 | 465 | 490 | 470 | 480 | 480 | 479.6184 | 450 | 472 | 495 | 516.6 | 478 | 541 | 486.45 | 477 | 476.93
| Std Dev. |  | 7.3139 | 3.5127 | 2.5543 | 7.0828 | 2.0656 | 5.4772 | 5.164 | 0 | 5.164 | 10.328 | 8.6858 | 3.7238 | 4.1191 | 12.2868 | 3.3194 | 6.442 | 9.347 | 5.7003 | 5.4314 | 3.0256
| Rel.Std.Dev. |  | 1.497% | 0.7361% | 0.5545% | 1.428% | 0.4128% | 1.178% | 1.061% | 0% | 1.083% | 2.137% | 1.82% | 0.8269% | 0.8761% | 2.463% | 0.6414% | 1.341% | 1.73% | 1.169% | 1.135% | 0.6338%
| PDM3 |  | 1.955% | -0.4079% | -3.856% | 3.482% | 4.421% | -2.953% | 1.569% | -1.91% | -0.5181% | 0.8732% | -0.3925% | -6.014% | -1.875% | 4.108% | 8.018% | 0.2819% | 12.73% | 1.746% | -0.1355% | -0.375%
| Z-Score (Absolute) |  | 0.66 | 0.17 | 1.38 | 1.2 | 1.53 | 1.07 | 0.53 | 0.7 | 0.21 | 0.28 | 0.16 | 2.14 | 0.69 | 1.42 | 2.8 | 0.07 | 4.46 | 0.59 | 0.07 | 0.16
| Analytical results for Be in OREAS 263b (Certified Value 2.16 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 2.35 | 2.1137 | 2.2827 | 2.9 | 2.1 | 2.16 | 2.13 | 2.23 | 2.21 | 2.03 | 2.3618 | 2.13 | 2.1 | 2 | 2.08 | 2 | 1.91 | 7.8 | 2.5 | 1.919 | 1
| 1 | 2 | 2.25 | 1.9861 | 2.3119 | 2.6 | 2 | 2.16 | 2.16 | 2.21 | 2.19 | 2.1 | 2.3565 | 2.14 | 2.1 | 2 | 2.14 | 2 | 2.29 | 7.68 | 2.5 | 1.959 | 25
| 1 | 3 | 2.13 | 2.1048 | 2.2971 | 2.8 | 2.2 | 2.17 | 2.15 | 2.22 | 2.2 | 2.05 | 2.4664 | 2.12 | 2.1 | 2 | 2.24 | 2 | 1.76 | 7.99 | 2.5 | 1.936 | 16
| 1 | 4 | 2.23 | 2.0036 | 2.2817 | 2.7 | 1.9 | 2.19 | 2.16 | 2.25 | 2.22 | 2.1 | 2.3684 | 2.12 | 2.1 | 2 | 2.21 | 2 | 2.24 | 7.74 | 2.5 | 1.919 | 2.1588
| 1 | 5 | 2.26 | 2.0616 | 2.3545 | 2.6 | 2.3 | 2.16 | 2.18 | 2.22 | 2.28 | 2.07 | 2.3499 | 2.14 | 2.1 | 2 | 2.14 | 2 | 2.43 | 7.81 | 2.5 | 1.876 | 19
| 1 | 6 | 2.18 | 1.9815 | 2.254 | 2.9 | 2.1 | 2.21 | 2.16 | 2.22 | 2.23 | 2.08 | 2.4657 | 2.14 | 2.1 | 2 | 2.06 | 2 | 2.4 | 7.55 | 2.5 | 1.895
| Mean |  | 2.2333 | 2.0419 | 2.297 | 2.75 | 2.1 | 2.175 | 2.1567 | 2.225 | 2.2217 | 2.0717 | 2.3948 | 2.1317 | 2.1 | 2 | 2.145 | 2 | 2.1717 | 7.7617 | 2.5 | 1.9173
| Median |  | 2.24 | 2.0326 | 2.2899 | 2.75 | 2.1 | 2.165 | 2.16 | 2.22 | 2.215 | 2.075 | 2.3651 | 2.135 | 2.1 | 2 | 2.14 | 2 | 2.265 | 7.77 | 2.5 | 1.919
| Std Dev. |  | 0.075011 | 0.05955 | 0.034119 | 0.13784 | 0.14142 | 0.020736 | 0.01633 | 0.013784 | 0.031885 | 0.027869 | 0.055526 | 0.0098319 | 0 | 0 | 0.070356 | 0 | 0.27404 | 0.14689 | 0 | 0.029317
| Rel.Std.Dev. |  | 3.359% | 2.916% | 1.485% | 5.012% | 6.734% | 0.9534% | 0.7572% | 0.6195% | 1.435% | 1.345% | 2.319% | 0.4612% | 0% | 0% | 3.28% | 0% | 12.62% | 1.893% | 0% | 1.529%
| PDM3 |  | 3.45% | -5.417% | 6.398% | 27.38% | -2.726% | 0.7484% | -0.1008% | 3.064% | 2.91% | -4.038% | 10.93% | -1.259% | -2.726% | -7.358% | -0.6412% | -7.358% | 0.594% | 259.5% | 15.8% | -11.19%
| Z-Score (Absolute) |  | 0.55 | 1.21 | 1.14 | 5.3 | 0.67 | 0.02 | 0.15 | 0.48 | 0.44 | 0.93 | 2.04 | 0.38 | 0.67 | NA | 0.26 | NA | 0.02 | 51.39 | 3 | 2.35
Results from laboratories 1.17 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 263b (Certified Value 10.1 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 9.5 | 9.4603 | 10.2 | 11.18 | 10.2 | 10 | 9.75 | 9.86 | 10.55 | 9.5684 | 10.4 | 10.3 | 8.91 | 10.46 | 10.09 | 8.95 | 6.87 | 11.4 | 4.329 |  | 1
| 1 | 2 | 9.99 | 9.4755 | 10.1 | 10.67 | 9.94 | 10.25 | 11 | 11.15 | 10.95 | 9.6552 | 10.22 | 10.5 | 8.83 | 10.2 | 9.64 | 9.11 | 6.37 | 11.3 | 3.317 |  | 26
| 1 | 3 | 10.33 | 10.017 | 9.77 | 11.28 | 10.6 | 10.2 | 10 | 9.77 | 11.2 | 9.6436 | 10.15 | 10.7 | 9.91 | 10.27 | 10.16 | 8.8 | 6.2 | 10.4 | 4.338 |  | 16
| 1 | 4 | 10.32 | 9.532 | 9.89 | 10.8 | 11.05 | 9.85 | 10.75 | 9.07 | 10.7 | 9.5813 | 9.74 | 10.1 | 9.16 | 10.43 | 10.21 | 9.93 | 6.25 | 10.5 | 3.867 |  | 10.1057
| 1 | 5 | 9.02 | 9.2717 | 10.3 | 10.41 | 9.91 | 9.13 | 10.15 | 9.46 | 11 | 9.6808 | 9.99 | 10.7 | 9.72 | 10.02 | 12.68 | 12.1 | 6.62 | 11.8 | 4.03 |  | 20
| 1 | 6 | 9.74 | 9.8926 | 9.79 | 10.9 | 10.15 | 10.35 | 10.3 | 9.95 | 11.6 | 9.6923 | 10.23 | 9.86 | 9.38 | 10.76 | 10.53 | 9.14 | 6.79 | 10.6 | 5.573
| Mean |  | 9.8167 | 9.6082 | 10.0083 | 10.8733 | 10.3083 | 9.9633 | 10.325 | 9.8767 | 11 | 9.6369 | 10.1217 | 10.36 | 9.3183 | 10.3567 | 10.5517 | 9.6717 | 6.5167 | 11 | 4.2423
| Median |  | 9.865 | 9.5037 | 9.995 | 10.85 | 10.175 | 10.1 | 10.225 | 9.815 | 10.975 | 9.6494 | 10.185 | 10.4 | 9.27 | 10.35 | 10.185 | 9.125 | 6.495 | 10.95 | 4.1795
| Std Dev. |  | 0.50749 | 0.28516 | 0.22302 | 0.32297 | 0.43961 | 0.44639 | 0.46984 | 0.70182 | 0.37283 | 0.051301 | 0.22921 | 0.33823 | 0.43485 | 0.25461 | 1.0812 | 1.2524 | 0.28395 | 0.57619 | 0.75227
| Rel.Std.Dev. |  | 5.17% | 2.968% | 2.228% | 2.97% | 4.265% | 4.48% | 4.551% | 7.106% | 3.389% | 0.5323% | 2.265% | 3.265% | 4.667% | 2.458% | 10.25% | 12.95% | 4.357% | 5.238% | 17.73%
| PDM3 |  | -2.86% | -4.923% | -0.9638% | 7.596% | 2.005% | -1.409% | 2.17% | -2.267% | 8.849% | -4.639% | 0.1577% | 2.516% | -7.792% | 2.483% | 4.413% | -4.295% | -35.52% | 8.849% | -58.02%
| Z-Score (Absolute) |  | 0.37 | 0.77 | 0 | 1.66 | 0.58 | 0.09 | 0.61 | 0.25 | 1.9 | 0.71 | 0.22 | 0.67 | 1.32 | 0.67 | 1.04 | 0.65 | 6.7 | 1.9 | 11.06
| Analytical results for Ca in OREAS 263b (Certified Value 0.294 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.2876 | 0.28468 | 0.29595 | 0.31 | 0.29 | 0.29 | 0.3 | 0.29 | 0.29 | 0.29 | 0.29847 | 0.3054 | 0.29 | 0.33 | 0.298 | 0.29 | 0.235 | 0.3 | 0.29 | 0.28214 | 1
| 1 | 2 | 0.2944 | 0.28725 | 0.29667 | 0.31 | 0.3 | 0.29 | 0.31 | 0.29 | 0.29 | 0.29 | 0.29517 | 0.3048 | 0.282 | 0.33 | 0.3031 | 0.28 | 0.23 | 0.3 | 0.3 | 0.28058 | #N/A
| 1 | 3 | 0.2893 | 0.28375 | 0.2969 | 0.32 | 0.29 | 0.3 | 0.31 | 0.29 | 0.29 | 0.28 | 0.29443 | 0.3022 | 0.286 | 0.29 | 0.2974 | 0.28 | 0.21 | 0.31 | 0.3 | 0.29685 | 16
| 1 | 4 | 0.2889 | 0.28226 | 0.2943 | 0.31 | 0.29 | 0.3 | 0.31 | 0.3 | 0.29 | 0.3 | 0.29648 | 0.3067 | 0.283 | 0.37 | 0.3046 | 0.28 | 0.23 | 0.3 | 0.29 | 0.2992 | 0.29448
| 1 | 5 | 0.2936 | 0.28825 | 0.29788 | 0.3 | 0.28 | 0.29 | 0.31 | 0.29 | 0.29 | 0.28 | 0.2982 | 0.3046 | 0.286 | 0.3 | 0.2993 | 0.28 | 0.256 | 0.3 | 0.3 | 0.27794 | 21
| 1 | 6 | 0.294 | 0.28429 | 0.29576 | 0.3 | 0.29 | 0.3 | 0.31 | 0.29 | 0.29 | 0.3 | 0.29813 | 0.3085 | 0.281 | 0.29 | 0.2959 | 0.29 | 0.224 | 0.3 | 0.3 | 0.2825
| Mean |  | 0.2913 | 0.28508 | 0.29624 | 0.30833 | 0.29 | 0.295 | 0.30833 | 0.29167 | 0.29 | 0.29 | 0.29682 | 0.30537 | 0.28467 | 0.31833 | 0.29972 | 0.28333 | 0.23083 | 0.30167 | 0.29667 | 0.28654
| Median |  | 0.29145 | 0.28449 | 0.29631 | 0.31 | 0.29 | 0.295 | 0.31 | 0.29 | 0.29 | 0.29 | 0.29731 | 0.3051 | 0.2845 | 0.315 | 0.29865 | 0.28 | 0.23 | 0.3 | 0.3 | 0.28232
| Std Dev. |  | 0.0030213 | 0.002248 | 0.0012133 | 0.0075277 | 0.0063246 | 0.0054772 | 0.0040825 | 0.0040825 | 0 | 0.0089443 | 0.0017238 | 0.0021229 | 0.0033267 | 0.031252 | 0.0034161 | 0.005164 | 0.015052 | 0.0040825 | 0.005164 | 0.0090742
| Rel.Std.Dev. |  | 1.037% | 0.7886% | 0.4096% | 2.441% | 2.181% | 1.857% | 1.324% | 1.4% | 0% | 3.084% | 0.5808% | 0.6952% | 1.169% | 9.817% | 1.14% | 1.823% | 6.521% | 1.353% | 1.741% | 3.167%
| PDM3 |  | -1.081% | -3.194% | 0.5972% | 4.703% | -1.523% | 0.175% | 4.703% | -0.9569% | -1.523% | -1.523% | 0.7914% | 3.695% | -3.334% | 8.098% | 1.777% | -3.787% | -21.61% | 2.439% | 0.741% | -2.699%
| Z-Score (Absolute) |  | 0.21 | 0.84 | 0.3 | 1.53 | 0.34 | 0.17 | 1.53 | 0.17 | 0.34 | 0.34 | 0.36 | 1.23 | 0.89 | 2.56 | 0.65 | 1.02 | 6.39 | 0.85 | 0.34 | 0.7
| Analytical results for Cd in OREAS 263b (Certified Value 0.044 ppm) | [21 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.05 | < 0.2 | < 0.3 | 0.22 | 0.03 | < 0.02 | < 0.02 | 0.03 | 0.02 | 0.081463 | 0.05 | < 0.05 | 0.07 | 0.04 | 0.04 | 0.06 | 2.34 | < 0.5 | 0.469 |  | 1
| 1 | 2 | 0.05 | < 0.2 | < 0.3 | 0.24 | 0.02 | 0.04 | 0.03 | < 0.02 | 0.03 | 0.084924 | 0.05 | < 0.05 | 0.06 | 0.06 | 0.03 | 0.07 | 2.13 | < 0.5 | 0.4 |  | 27
| 1 | 3 | 0.03 | < 0.2 | < 0.3 | 0.23 | 0.02 | 0.04 | < 0.02 | 0.02 | 0.02 | 0.064261 | 0.05 | < 0.05 | 0.04 | 0.04 | 0.04 | < 0.05 | 2.39 | < 0.5 | 0.502 |  | 16
| 1 | 4 | 0.03 | < 0.2 | < 0.3 | 0.21 | 0.03 | 0.03 | < 0.02 | 0.03 | < 0.02 | 0.051355 | 0.05 | < 0.05 | 0.11 | 0.04 | 0.05 | 0.05 | 2.17 | < 0.5 | 0.383 |  | 0.04369
| 1 | 5 | 0.04 | < 0.2 | < 0.3 | 0.24 | 0.02 | < 0.02 | 0.08 | 0.02 | 0.03 | 0.045882 | 0.05 | < 0.05 | 0.07 | 0.04 | 0.03 | 0.05 | 2.2 | < 0.5 | 0.475 |  | 22
| 1 | 6 | 0.05 | < 0.2 | < 0.3 | 0.24 | 0.02 | 0.03 | 0.04 | < 0.02 | 0.04 | 0.055779 | 0.06 | < 0.05 | 0.09 | 0.05 | 0.03 | 0.06 | 2.13 | < 0.5 | 0.405
| Mean |  | 0.041667 |  |  | 0.23 | 0.023333 | 0.035 | 0.05 | 0.025 | 0.028 | 0.063944 | 0.051667 |  | 0.073333 | 0.045 | 0.036667 | 0.058 | 2.2267 |  | 0.439
| Median |  | 0.045 |  |  | 0.235 | 0.02 | 0.035 | 0.04 | 0.025 | 0.03 | 0.06002 | 0.05 |  | 0.07 | 0.04 | 0.035 | 0.06 | 2.185 |  | 0.437
| Std Dev. |  | 0.0098319 |  |  | 0.012649 | 0.005164 | 0.0057735 | 0.026458 | 0.0057735 | 0.0083666 | 0.016116 | 0.0040825 |  | 0.024221 | 0.0083666 | 0.008165 | 0.0083666 | 0.11147 |  | 0.048945
| Rel.Std.Dev. |  | 23.6% |  |  | 5.5% | 22.13% | 16.5% | 52.92% | 23.09% | 29.88% | 25.2% | 7.902% |  | 33.03% | 18.59% | 22.27% | 14.43% | 5.006% |  | 11.15%
| PDM3 |  | -4.631% |  |  | 426.4% | -46.59% | -19.89% | 14.44% | -42.78% | -35.91% | 46.36% | 18.26% |  | 67.85% | 2.999% | -16.07% | 32.75% | 4997% |  | 904.8%
| Z-Score (Absolute) |  | 0.25 | 1.22 | 2.49 | 4.51 | 0.72 | 0.63 | 0.55 | 0.8 | 0.67 | 0.31 | 0 | 0.67 | 0.55 | 0.17 | 0.38 | 0.02 | 55 | 5.02 | 9.79
| Analytical results for Ce in OREAS 263b (Certified Value 82 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 0
1
| 1 | 1 | 82.98 | 80.8526 | 85.3705 | 84.5 | 86.6 | 82 | 82.5 | 80.2 | 81.9 | 67.4 | 66.36 | 85.48 | 78.96 | 79.1 | 82.2 | 74.328 |  |  |  |  | 1
| 1 | 2 | 83.06 | 79.5918 | 85.2652 | 83.7 | 86.9 | 88.9 | 81.8 | 79.9 | 81.3 | 66.2 | 61.75 | 88.05 | 79.08 | 81.7 | 85 | 75.811 |  |  |  |  | 28
| 1 | 3 | 82.78 | 80.6343 | 85.2619 | 84.8 | 87.7 | 80.9 | 81.6 | 75.3 | 90 | 65.3 | 71.61 | 87.25 | 77.21 | 78.4 | 83.6 | 71.873 |  |  |  |  | 16
| 1 | 4 | 83.09 | 82.2406 | 85.261 | 83.8 | 88 | 81.6 | 83.4 | 77.4 | 84.8 | 65.7 | 65.03 | 85.94 | 80.06 | 77.7 | 83.6 | 72.178 |  |  |  |  | 82.2028
| 1 | 5 | 82.66 | 81.4271 | 85.2889 | 83.7 | 86.3 | 76.2 | 81.8 | 77.8 | 81.8 | 64.9 | 80.74 | 87.87 | 80.34 | 80 | 84.9 | 77.36 |  |  |  |  | 23
| 1 | 6 | 82.47 | 82.5126 | 85.2328 | 83.4 | 86.4 | 84.8 | 84.6 | 79.6 | 84.1 | 66 | 72.02 | 88.5 | 80.54 | 77.3 | 84.9 | 77.048
| Mean |  | 82.84 | 81.2098 | 85.2801 | 83.9833 | 86.9833 | 82.4 | 82.6167 | 78.3667 | 83.9833 | 65.9167 | 69.585 | 87.1817 | 79.365 | 79.0333 | 84.0333 | 74.7663
| Median |  | 82.88 | 81.1398 | 85.2636 | 83.75 | 86.75 | 81.8 | 82.15 | 78.7 | 83 | 65.85 | 68.985 | 87.56 | 79.57 | 78.75 | 84.25 | 75.0695
| Std Dev. |  | 0.24633 | 1.0848 | 0.047736 | 0.54191 | 0.70828 | 4.2308 | 1.1771 | 1.8917 | 3.2615 | 0.86583 | 6.7379 | 1.2172 | 1.2413 | 1.6269 | 1.1112 | 2.3785
| Rel.Std.Dev. |  | 0.2974% | 1.336% | 0.05598% | 0.6453% | 0.8143% | 5.135% | 1.425% | 2.414% | 3.884% | 1.314% | 9.683% | 1.396% | 1.564% | 2.058% | 1.322% | 3.181%
| PDM3 |  | 0.7751% | -1.208% | 3.743% | 2.166% | 5.816% | 0.2399% | 0.5034% | -4.667% | 2.166% | -19.81% | -15.35% | 6.057% | -3.452% | -3.856% | 2.227% | -9.047%
| Z-Score (Absolute) |  | 0.08 | 0.3 | 0.63 | 0.34 | 1.02 | 0.02 | 0.02 | 0.94 | 0.34 | 3.78 | 2.95 | 1.07 | 0.72 | 0.79 | 0.35 | 1.77
| Analytical results for Co in OREAS 263b (Certified Value 4.81 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 5 | 4.4274 | 4.8443 | 5 | 4.8 | 4.9 | 4.8 | 4.6 | 4.5 | 4.5 | 4.6787 | 4.9 | 5 | 4.7 | 5 | 4.6 | 4.3 | 5.03 | 5 | 5.305 | 1
| 1 | 2 | 5 | 4.2695 | 5.3699 | 5 | 4.8 | 4.8 | 4.9 | 4.7 | 4.7 | 4.6 | 4.5277 | 4.8 | 5 | 4.6 | 5 | 4.6 | 4.6 | 5.04 | 5 | 5.034 | 14
| 1 | 3 | 5.3 | 4.335 | 4.9316 | 5 | 4.6 | 4.9 | 4.5 | 4.6 | 4.7 | 4.9 | 4.5379 | 4.9 | 5 | 4.7 | 5 | 4.6 | 4.4 | 5.04 | 5 | 5.432 | 16
| 1 | 4 | 5 | 4.3668 | 5.289 | 5 | 5.1 | 4.9 | 4.6 | 4.7 | 4.6 | 5.1 | 4.8948 | 4.9 | 5 | 4.9 | 5.1 | 4.6 | 4.6 | 5.08 | 5 | 4.979 | 4.8092
| 1 | 5 | 5.1 | 4.4438 | 4.583 | 5 | 4.8 | 4.9 | 4.4 | 4.8 | 4.5 | 4.7 | 4.6773 | 5 | 5 | 4.9 | 5.1 | 4.7 | 5 | 5.06 | 5 | 5.53 | 24
| 1 | 6 | 5.1 | 4.4003 | 4.9251 | 5 | 4.5 | 4.9 | 4.8 | 4.8 | 4.7 | 4.9 | 4.8371 | 5.1 | 5 | 4.8 | 5.1 | 4.7 | 4.4 | 5.09 | 5 | 4.975
| Mean |  | 5.0833 | 4.3738 | 4.9905 | 5 | 4.7667 | 4.8833 | 4.6667 | 4.7 | 4.6167 | 4.7833 | 4.6922 | 4.9333 | 5 | 4.7667 | 5.05 | 4.6333 | 4.55 | 5.0567 | 5 | 5.2092
| Median |  | 5.05 | 4.3835 | 4.9284 | 5 | 4.8 | 4.9 | 4.7 | 4.7 | 4.65 | 4.8 | 4.678 | 4.9 | 5 | 4.75 | 5.05 | 4.6 | 4.5 | 5.05 | 5 | 5.1695
| Std Dev. |  | 0.1169 | 0.06466 | 0.29263 | 0 | 0.20656 | 0.040825 | 0.19664 | 0.089443 | 0.098319 | 0.22286 | 0.15054 | 0.10328 | 0 | 0.12111 | 0.054772 | 0.05164 | 0.251 | 0.024221 | 0 | 0.24506
| Rel.Std.Dev. |  | 2.3% | 1.478% | 5.864% | 0% | 4.333% | 0.836% | 4.214% | 1.903% | 2.13% | 4.659% | 3.208% | 2.094% | 0% | 2.541% | 1.085% | 1.115% | 5.516% | 0.479% | 0% | 4.704%
| PDM3 |  | 5.701% | -9.053% | 3.771% | 3.968% | -0.8835% | 1.542% | -2.963% | -2.27% | -4.003% | -0.537% | -2.431% | 2.582% | 3.968% | -0.8835% | 5.008% | -3.656% | -5.389% | 5.147% | 3.968% | 8.318%
| Z-Score (Absolute) |  | 1.42 | 1.77 | 1.01 | NA | 0 | 0.52 | 0.45 | 0.3 | 0.67 | 0.07 | 0.33 | 0.75 | NA | 0 | 1.27 | 0.6 | 0.97 | 1.3 | NA | 1.99
| Analytical results for Cr in OREAS 263b (Certified Value 78 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 0
1
| 1 | 1 | 85 | 77.2697 | 50 | 77 | 77 | 79 | 76 | 79 | 78 | 74.9746 | 81.5 | 84 | 60 | 88 | 65 | 71 | 80.2 | 60 | 83.529 |  | 1
| 1 | 2 | 83 | 75.7947 | 44 | 74 | 77 | 77 | 76 | 78 | 79 | 73.7609 | 82 | 85 | 58 | 85 | 63 | 67 | 81.4 | 80 | 90.714 |  | 30
| 1 | 3 | 85 | 74.8476 | 47 | 75 | 81 | 77 | 75 | 78 | 77 | 73.5273 | 80.2 | 87 | 62 | 81 | 66 | 69 | 80.6 | 60 | 89.236 |  | 16
| 1 | 4 | 84 | 76.5283 | 45 | 73 | 81 | 79 | 78 | 79 | 80 | 75.5151 | 81.3 | 83 | 74 | 87 | 65 | 70 | 83.4 | 80 | 89.106 |  | 78.1309
| 1 | 5 | 89 | 76.7153 | 50 | 74 | 76 | 76 | 77 | 80 | 75 | 74.8451 | 81.7 | 86 | 65 | 85 | 72 | 73 | 79.03 | 80 | 91.448 |  | 25
| 1 | 6 | 85 | 77.2063 | 53 | 75 | 80 | 78 | 76 | 78 | 78 | 74.8451 | 82.3 | 83 | 62 | 77 | 66 | 69 | 78.1 | 80 | 91.492
| Mean |  | 85.1667 | 76.3937 | 48.1667 | 74.6667 | 78.6667 | 77.6667 | 76.3333 | 78.6667 | 77.8333 | 74.578 | 81.5 | 84.6667 | 63.5 | 83.8333 | 66.1667 | 69.8333 | 80.455 | 73.3333 | 89.2542
| Median |  | 85 | 76.6218 | 48.5 | 74.5 | 78.5 | 77.5 | 76 | 78.5 | 78 | 74.8451 | 81.6 | 84.5 | 62 | 85 | 65.5 | 69.5 | 80.4 | 80 | 89.975
| Std Dev. |  | 2.0412 | 0.92719 | 3.4303 | 1.3663 | 2.2509 | 1.2111 | 1.0328 | 0.8165 | 1.7224 | 0.76809 | 0.72938 | 1.633 | 5.648 | 4.1191 | 3.0605 | 2.0412 | 1.8563 | 10.328 | 2.9917
| Rel.Std.Dev. |  | 2.397% | 1.214% | 7.122% | 1.83% | 2.861% | 1.559% | 1.353% | 1.038% | 2.213% | 1.03% | 0.8949% | 1.929% | 8.895% | 4.913% | 4.625% | 2.923% | 2.307% | 14.08% | 3.352%
| PDM3 |  | 9.005% | -2.224% | -38.35% | -4.434% | 0.6857% | -0.5942% | -2.301% | 0.6857% | -0.3809% | -4.547% | 4.312% | 8.365% | -18.73% | 7.299% | -15.31% | -10.62% | 2.975% | -6.14% | 14.24%
| Z-Score (Absolute) |  | 1.32 | 0.22 | 5.19 | 0.53 | 0.18 | 0 | 0.23 | 0.18 | 0.03 | 0.54 | 0.67 | 1.23 | 2.49 | 1.08 | 2.02 | 1.38 | 0.49 | 0.76 | 2.04
| Analytical results for Cs in OREAS 263b (Certified Value 8.01 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 7.95 | 8.1465 | 8.01 | 8.44 | 8.09 | 7.51 | 7.81 | 8.18 | 8.092 | 7.9 | 7.8 | 7.4 | 8.5 | 8.3 | 9.3 | 7.9 |  |  |  |  | 1
| 1 | 2 | 8 | 8.045 | 8.64 | 8.52 | 8.11 | 7.59 | 7.6 | 8.31 | 8.0056 | 7.7 | 7.9 | 7.4 | 8.57 | 7.9 | 9.3 | 8 |  |  |  |  | 31
| 1 | 3 | 8.09 | 8.0273 | 8.13 | 8.57 | 7.72 | 7.53 | 7.56 | 9.03 | 7.9003 | 7.8 | 7.9 | 7.6 | 8.5 | 7.7 | 9.1 | 7.9 |  |  |  |  | 16
| 1 | 4 | 8 | 8.2032 | 8.22 | 8.5 | 8.03 | 7.75 | 7.62 | 8.64 | 8.098 | 7.7 | 7.8 | 7.2 | 8.5 | 8.1 | 8.9 | 8 |  |  |  |  | 8.0141
| 1 | 5 | 7.85 | 8.0537 | 8.19 | 8.59 | 7.28 | 7.57 | 7.33 | 8.52 | 8.1685 | 7.8 | 7.9 | 8 | 8.54 | 8.3 | 9.5 | 8.1 |  |  |  |  | 26
| 1 | 6 | 7.86 | 7.9772 | 8.2 | 8.61 | 8.12 | 7.78 | 7.46 | 8.68 | 7.865 | 7.9 | 7.9 | 7.1 | 8.54 | 8.2 | 9.1 | 8
| Mean |  | 7.9583 | 8.0755 | 8.2317 | 8.5383 | 7.8917 | 7.6217 | 7.5633 | 8.56 | 8.0216 | 7.8 | 7.8667 | 7.45 | 8.525 | 8.0833 | 9.2 | 7.9833
| Median |  | 7.975 | 8.0493 | 8.195 | 8.545 | 8.06 | 7.58 | 7.58 | 8.58 | 8.0488 | 7.8 | 7.9 | 7.4 | 8.52 | 8.15 | 9.2 | 8
| Std Dev. |  | 0.09196 | 0.083335 | 0.21405 | 0.063692 | 0.33523 | 0.11496 | 0.16158 | 0.30007 | 0.1199 | 0.089443 | 0.05164 | 0.32094 | 0.029496 | 0.24014 | 0.20976 | 0.075277
| Rel.Std.Dev. |  | 1.156% | 1.032% | 2.6% | 0.746% | 4.248% | 1.508% | 2.136% | 3.505% | 1.495% | 1.147% | 0.6564% | 4.308% | 0.346% | 2.971% | 2.28% | 0.9429%
| PDM3 |  | -0.6955% | 0.7662% | 2.715% | 6.542% | -1.527% | -4.896% | -5.624% | 6.812% | 0.09352% | -2.671% | -1.839% | -7.038% | 6.375% | 0.8643% | 14.8% | -0.3835%
| Z-Score (Absolute) |  | 0.14 | 0.23 | 0.72 | 1.67 | 0.35 | 1.19 | 1.37 | 1.74 | 0.06 | 0.63 | 0.42 | 1.73 | 1.63 | 0.25 | 3.74 | 0.06
| Analytical results for Cu in OREAS 263b (Certified Value 75 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 76.2 | 72.5367 | 72.823 | 79 | 74 | 71.9 | 75.8 | 73.7 | 74.5 | 72.9 | 69.5584 | 68.6 | 71 | 74.8 | 76.3 | 79.7 | 72.7 | 78.26 | 64 | 78.808 | 1
| 1 | 2 | 77.4 | 71.1088 | 74.8075 | 78 | 75 | 71.3 | 77.1 | 73.5 | 74 | 73.8 | 71.1073 | 67.2 | 71 | 73 | 78 | 75.7 | 68.8 | 78.08 | 64 | 81.358 | 16
| 1 | 3 | 79.2 | 71.5599 | 75.8155 | 80 | 81 | 73 | 74 | 75.9 | 73.5 | 78.8 | 68.0384 | 65.7 | 73 | 77.4 | 79.9 | 77.4 | 62.4 | 79.55 | 60 | 79.586 | 16
| 1 | 4 | 79.7 | 71.6582 | 73.684 | 77 | 78 | 74.3 | 75.3 | 76.3 | 71.4 | 78.2 | 69.7875 | 66.8 | 74 | 77.9 | 75.8 | 77.3 | 72.8 | 78.63 | 64 | 78.775 | 74.9709
| 1 | 5 | 76.3 | 72.3444 | 73.9885 | 93 | 76 | 72.1 | 72.1 | 75.1 | 77.2 | 75 | 71.6112 | 72 | 74 | 81.3 | 77.3 | 79.5 | 76.3 | 80.12 | 64 | 80.404 | 27
| 1 | 6 | 73.9 | 71.6774 | 75.0175 | 75 | 76 | 72.3 | 75.7 | 75.3 | 73.1 | 75.5 | 69.2819 | 69 | 72 | 78.1 | 75.5 | 82.4 | 70.8 | 78.55 | 62 | 80.771
| Mean |  | 77.1167 | 71.8142 | 74.356 | 80.3333 | 76.6667 | 72.4833 | 75 | 74.9667 | 73.95 | 75.7 | 69.8974 | 68.2167 | 72.5 | 77.0833 | 77.1333 | 78.6667 | 70.6333 | 78.865 | 63 | 79.9503
| Median |  | 76.85 | 71.6678 | 74.398 | 78.5 | 76 | 72.2 | 75.5 | 75.2 | 73.75 | 75.25 | 69.6729 | 67.9 | 72.5 | 77.65 | 76.8 | 78.45 | 71.75 | 78.59 | 64 | 79.995
| Std Dev. |  | 2.1424 | 0.53073 | 1.0685 | 6.4395 | 2.5033 | 1.0477 | 1.7344 | 1.1431 | 1.9129 | 2.3597 | 1.2933 | 2.2113 | 1.3784 | 2.8799 | 1.6476 | 2.3653 | 4.7382 | 0.79762 | 1.6733 | 1.0658
| Rel.Std.Dev. |  | 2.778% | 0.739% | 1.437% | 8.016% | 3.265% | 1.445% | 2.312% | 1.525% | 2.587% | 3.117% | 1.85% | 3.242% | 1.901% | 3.736% | 2.136% | 3.007% | 6.708% | 1.011% | 2.656% | 1.333%
| PDM3 |  | 2.862% | -4.21% | -0.8201% | 7.153% | 2.262% | -3.318% | 0.03887% | -0.005593% | -1.362% | 0.9726% | -6.767% | -9.009% | -3.296% | 2.818% | 2.884% | 4.93% | -5.786% | 5.194% | -15.97% | 6.642%
| Z-Score (Absolute) |  | 0.58 | 0.86 | 0.17 | 1.45 | 0.46 | 0.68 | 0 | 0 | 0.28 | 0.19 | 1.38 | 1.83 | 0.67 | 0.57 | 0.58 | 1 | 1.18 | 1.05 | 3.24 | 1.34
| Analytical results for Dy in OREAS 263b (Certified Value 3.21 ppm) | [21 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 3.2 | 3.2189 | 3.3273 | 3.8 | 2.8 | 3.34 | 2.9 | 3.02 | 3.5 | 3.224 | [10 empty] | 1
| 1 | 2 | 3.14 | 2.9506 | 3.2843 | 3.6 | 2.6 | 3.46 | 2.9 | 3.18 | 3.6 | 3.296 | [10 empty] | 4
| 1 | 3 | 3.16 | 2.9661 | 3.2795 | 3.6 | 2.9 | 3.39 | 2.9 | 3 | 3.6 | 3.01 | [10 empty] | 16
| 1 | 4 | 3.22 | 3.0674 | 3.3807 | 3.7 | 2.6 | 3.34 | 3 | 2.94 | 3.5 | 3.01 | [10 empty] | 3.2126
| 1 | 5 | 3.23 | 2.997 | 3.3565 | 3.6 | 3.2 | 3.96 | 3 | 3.65 | 3.55 | 3.332 | [10 empty] | 28
| 1 | 6 | 3.16 | 3.0726 | 3.3892 | 3.8 | 2.8 | 3.4 | 2.9 | 2.96 | 3.5 | 3.197
| Mean |  | 3.185 | 3.0454 | 3.3362 | 3.6833 | 2.8167 | 3.4817 | 2.9333 | 3.125 | 3.5417 | 3.1782
| Median |  | 3.18 | 3.0322 | 3.3419 | 3.65 | 2.8 | 3.395 | 2.9 | 3.01 | 3.525 | 3.2105
| Std Dev. |  | 0.036742 | 0.098921 | 0.047325 | 0.098319 | 0.22286 | 0.23853 | 0.05164 | 0.27083 | 0.04916 | 0.13897
| Rel.Std.Dev. |  | 1.154% | 3.248% | 1.418% | 2.669% | 7.912% | 6.851% | 1.76% | 8.667% | 1.388% | 4.373%
| PDM3 |  | -0.8586% | -5.203% | 3.849% | 14.65% | -12.32% | 8.376% | -8.692% | -2.726% | 10.24% | -1.071%
| Z-Score (Absolute) |  | 0.01 | 0.46 | 0.52 | 1.68 | 1.22 | 1 | 0.83 | 0.19 | 1.21 | 0.01
| Analytical results for Er in OREAS 263b (Certified Value 1.67 ppm) | [21 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 1.69 | 1.6093 | 1.745 | 2 | 1.4 | 1.66 | 1.4 | 1.5 | 1.8 | [11 empty] | 1
| 1 | 2 | 1.56 | 1.4952 | 1.8019 | 2 | 1.3 | 1.9 | 1.4 | 1.58 | 1.85 | [11 empty] | 5
| 1 | 3 | 1.62 | 1.4915 | 1.7742 | 2 | 1.5 | 1.7 | 1.4 | 1.54 | 1.9 | [11 empty] | 16
| 1 | 4 | 1.59 | 1.5639 | 1.7464 | 2 | 1.4 | 1.85 | 1.4 | 1.55 | 1.8 | [11 empty] | 1.6664
| 1 | 5 | 1.66 | 1.5765 | 1.753 | 2 | 1.6 | 1.81 | 1.4 | 1.87 | 1.8 | [11 empty] | 29
| 1 | 6 | 1.62 | 1.593 | 1.7479 | 2.1 | 1.4 | 1.99 | 1.4 | 1.57 | 1.9
| Mean |  | 1.6233 | 1.5549 | 1.7614 | 2.0167 | 1.4333 | 1.8183 | 1.4 | 1.6017 | 1.8417
| Median |  | 1.62 | 1.5702 | 1.7504 | 2 | 1.4 | 1.83 | 1.4 | 1.56 | 1.825
| Std Dev. |  | 0.046762 | 0.05006 | 0.022582 | 0.040825 | 0.10328 | 0.12352 | 2.4324e-16 | 0.13438 | 0.04916
| Rel.Std.Dev. |  | 2.881% | 3.219% | 1.282% | 2.024% | 7.206% | 6.793% | 1.737e-14% | 8.39% | 2.669%
| PDM3 |  | -2.585% | -6.691% | 5.701% | 21.02% | -13.99% | 9.117% | -15.99% | -3.885% | 10.52%
| Z-Score (Absolute) |  | 0 | 0.24 | 0.49 | 1.4 | 0.67 | 0.69 | 0.79 | 0.08 | 0.77
| Analytical results for Eu in OREAS 263b (Certified Value 1.02 ppm) | [21 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 1.05 | 1.0644 | 1.2981 | 1.07 | 0.9 | 1.08 | 1 | 0.91 | 1.2 | 0.96 | [10 empty] | 1
| 1 | 2 | 1.02 | 0.9976 | 1.3539 | 1.13 | 0.8 | 1.09 | 1 | 0.98 | 1.15 | 1.018 | [10 empty] | 35
| 1 | 3 | 1.06 | 0.99167 | 1.3543 | 1.09 | 1 | 1.06 | 0.9 | 0.92 | 1.15 | 0.944 | [10 empty] | 16
| 1 | 4 | 1.05 | 1.0028 | 1.3608 | 1.09 | 0.9 | 1.05 | 1 | 0.94 | 1.15 | 0.948 | [10 empty] | 1.0244
| 1 | 5 | 1.09 | 1.0435 | 1.3478 | 1.07 | 1.1 | 1 | 1 | 0.9 | 1.2 | 1.03 | [10 empty] | 30
| 1 | 6 | 1.02 | 1.0555 | 1.377 | 1.01 | 1 | 1.07 | 1 | 0.93 | 1.15 | 0.98
| Mean |  | 1.0483 | 1.0259 | 1.3487 | 1.0767 | 0.95 | 1.0583 | 0.98333 | 0.93 | 1.1667 | 0.98
| Median |  | 1.05 | 1.0232 | 1.3541 | 1.08 | 0.95 | 1.065 | 1 | 0.925 | 1.15 | 0.97
| Std Dev. |  | 0.026394 | 0.032144 | 0.026689 | 0.039328 | 0.10488 | 0.031885 | 0.040825 | 0.028284 | 0.02582 | 0.036508
| Rel.Std.Dev. |  | 2.518% | 3.133% | 1.979% | 3.653% | 11.04% | 3.013% | 4.152% | 3.041% | 2.213% | 3.725%
| PDM3 |  | 2.34% | 0.1503% | 31.66% | 5.106% | -7.259% | 3.317% | -4.005% | -9.212% | 13.89% | -4.33%
| Z-Score (Absolute) |  | 0.14 | 0.14 | 3.79 | 0.48 | 1.06 | 0.26 | 0.65 | 1.3 | 1.58 | 0.69
| Analytical results for Fe in OREAS 263b (Certified Value 2.68 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.63 | 2.5216 | 2.654 | 2.7 | 2.57 | 2.57 | 2.61 | 2.72 | 2.62 | 2.57 | 2.6494 | 2.6176 | 2.72 | 2.79 | 2.6998 | 2.72 | 3.2 | 2.78 | 2.6 | 2.9062 | 1
| 1 | 2 | 2.67 | 2.5318 | 2.6654 | 2.7 | 2.62 | 2.57 | 2.67 | 2.71 | 2.61 | 2.61 | 2.8597 | 2.6157 | 2.7 | 2.86 | 2.7065 | 2.71 | 3.07 | 2.77 | 2.6 | 2.9065 | #N/A
| 1 | 3 | 2.64 | 2.542 | 2.6623 | 2.73 | 2.63 | 2.59 | 2.64 | 2.71 | 2.6 | 2.54 | 2.6148 | 2.5739 | 2.74 | 2.75 | 2.6908 | 2.71 | 3.33 | 2.78 | 2.58 | 2.8755 | 16
| 1 | 4 | 2.63 | 2.5657 | 2.6693 | 2.71 | 2.56 | 2.61 | 2.69 | 2.76 | 2.64 | 2.65 | 2.8882 | 2.5844 | 2.7 | 3.03 | 2.6899 | 2.67 | 3.15 | 2.79 | 2.51 | 2.9019 | 2.6821
| 1 | 5 | 2.66 | 2.5781 | 2.6547 | 2.67 | 2.56 | 2.57 | 2.69 | 2.72 | 2.65 | 2.57 | 2.9308 | 2.5801 | 2.74 | 2.89 | 2.7227 | 2.74 | 2.65 | 2.73 | 2.6 | 2.9153 | 31
| 1 | 6 | 2.67 | 2.5581 | 2.6532 | 2.74 | 2.52 | 2.61 | 2.66 | 2.71 | 2.65 | 2.64 | 2.8859 | 2.6319 | 2.71 | 2.81 | 2.707 | 2.75 | 2.85 | 2.74 | 2.6 | 2.8861
| Mean |  | 2.65 | 2.5495 | 2.6598 | 2.7083 | 2.5767 | 2.5867 | 2.66 | 2.7217 | 2.6283 | 2.5967 | 2.8048 | 2.6006 | 2.7183 | 2.855 | 2.7028 | 2.7167 | 3.0417 | 2.765 | 2.5817 | 2.8986
| Median |  | 2.65 | 2.55 | 2.6585 | 2.705 | 2.565 | 2.58 | 2.665 | 2.715 | 2.63 | 2.59 | 2.8728 | 2.6 | 2.715 | 2.835 | 2.7031 | 2.715 | 3.11 | 2.775 | 2.6 | 2.904
| Std Dev. |  | 0.018974 | 0.021439 | 0.0068102 | 0.024833 | 0.041312 | 0.019664 | 0.030984 | 0.019408 | 0.02137 | 0.043665 | 0.13616 | 0.024051 | 0.018348 | 0.099146 | 0.012221 | 0.028048 | 0.24935 | 0.02429 | 0.036009 | 0.01483
| Rel.Std.Dev. |  | 0.716% | 0.8409% | 0.256% | 0.9169% | 1.603% | 0.7602% | 1.165% | 0.7131% | 0.8131% | 1.682% | 4.854% | 0.9248% | 0.675% | 3.473% | 0.4522% | 1.032% | 8.198% | 0.8785% | 1.395% | 0.5116%
| PDM3 |  | -1.198% | -4.944% | -0.8315% | 0.977% | -3.932% | -3.559% | -0.8251% | 1.474% | -2.006% | -3.186% | 4.574% | -3.04% | 1.35% | 6.445% | 0.77% | 1.288% | 13.4% | 3.09% | -3.746% | 8.07%
| Z-Score (Absolute) |  | 0.26 | 1.08 | 0.18 | 0.22 | 0.86 | 0.78 | 0.18 | 0.33 | 0.44 | 0.7 | 1.01 | 0.66 | 0.3 | 1.42 | 0.18 | 0.29 | 2.96 | 0.69 | 0.82 | 1.78
| Analytical results for Ga in OREAS 263b (Certified Value 16.6 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 17 | 16.7202 | 21.397 | 17.5 | 16.7 | 16.45 | 16.4 | 16.05 | 16.65 | 16.05 | 16.5202 | 17.21 | 17.5 | 15.14 | 18.25 | 15.66 | 16.1 | 16.2 |  |  | 1
| 1 | 2 | 17.32 | 16.5618 | 21.312 | 17.2 | 16.6 | 16.3 | 16.55 | 16.1 | 16.4 | 15.95 | 16.1367 | 16.91 | 17.2 | 14.93 | 18.45 | 15.67 | 16.6 | 16.6 |  |  | 36
| 1 | 3 | 17.15 | 16.2736 | 21.489 | 17.7 | 16.1 | 16.45 | 15.9 | 16.1 | 16.25 | 17.3 | 16.444 | 17.26 | 17.9 | 16.97 | 18.05 | 15.67 | 15.9 | 16.2 |  |  | 16
| 1 | 4 | 17.04 | 16.9074 | 21.397 | 17.8 | 16.5 | 16.5 | 16.15 | 16.65 | 15.9 | 15.8 | 16.9723 | 17.03 | 16.9 | 15.79 | 18.11 | 15.65 | 16.8 | 16.4 |  |  | 16.5685
| 1 | 5 | 16.97 | 16.8386 | 21.391 | 17.3 | 16 | 16.6 | 15.1 | 16.25 | 15.7 | 16.3 | 16.4482 | 17.05 | 16.8 | 17.56 | 18.7 | 16.17 | 18.2 | 16.6 |  |  | 32
| 1 | 6 | 17.03 | 16.1778 | 21.31 | 17.4 | 16.4 | 16.5 | 16.5 | 16.65 | 16.3 | 16.25 | 17.046 | 17.43 | 17.3 | 17.47 | 18.05 | 16.06 | 17 | 16.4
| Mean |  | 17.085 | 16.5799 | 21.3827 | 17.4833 | 16.3833 | 16.4667 | 16.1 | 16.3 | 16.2 | 16.275 | 16.5946 | 17.1483 | 17.2667 | 16.31 | 18.2683 | 15.8133 | 16.7667 | 16.4
| Median |  | 17.035 | 16.641 | 21.394 | 17.45 | 16.45 | 16.475 | 16.275 | 16.175 | 16.275 | 16.15 | 16.4842 | 17.13 | 17.25 | 16.38 | 18.18 | 15.67 | 16.7 | 16.4
| Std Dev. |  | 0.13035 | 0.29988 | 0.066425 | 0.23166 | 0.27869 | 0.098319 | 0.54681 | 0.27928 | 0.34496 | 0.53549 | 0.34806 | 0.18766 | 0.40332 | 1.1737 | 0.2608 | 0.23636 | 0.8165 | 0.17889
| Rel.Std.Dev. |  | 0.7629% | 1.809% | 0.3106% | 1.325% | 1.701% | 0.5971% | 3.396% | 1.713% | 2.129% | 3.29% | 2.097% | 1.094% | 2.336% | 7.196% | 1.428% | 1.495% | 4.87% | 1.091%
| PDM3 |  | 3.117% | 0.06854% | 29.06% | 5.521% | -1.118% | -0.6148% | -2.828% | -1.621% | -2.224% | -1.772% | 0.1572% | 3.499% | 4.214% | -1.56% | 10.26% | -4.558% | 1.196% | -1.017%
| Z-Score (Absolute) |  | 1.33 | 0.13 | 11.47 | 2.27 | 0.33 | 0.13 | 1 | 0.53 | 0.76 | 0.59 | 0.17 | 1.47 | 1.75 | 0.5 | 4.12 | 1.68 | 0.57 | 0.29
| Analytical results for Gd in OREAS 263b (Certified Value 4.26 ppm) | [21 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 4.19 | 4.31 | 4.8 | 4 | 4.43 | 4 | 4.08 | 4.2 | [12 empty] | 1
| 1 | 2 | 4.03 | 4.1036 | 4.8 | 3.8 | 4.38 | 4 | 4.22 | 4.4 | [12 empty] | 6
| 1 | 3 | 4.36 | 4.2021 | 4.6 | 4 | 4.39 | 3.8 | 4.24 | 4.2 | [12 empty] | 16
| 1 | 4 | 4.47 | 4.2087 | 4.7 | 3.8 | 4.32 | 4 | 4.11 | 4.2 | [12 empty] | 4.2644
| 1 | 5 | 4.26 | 4.4484 | 4.5 | 4.4 | 4.47 | 4 | 4.32 | 4.4 | [12 empty] | 33
| 1 | 6 | 4.59 | 4.3979 | 4.4 | 4 | 4.46 | 4.1 | 4.2 | 4.4
| Mean |  | 4.3167 | 4.2784 | 4.6333 | 4 | 4.4083 | 3.9833 | 4.195 | 4.3
| Median |  | 4.31 | 4.2594 | 4.65 | 4 | 4.41 | 4 | 4.21 | 4.3
| Std Dev. |  | 0.20077 | 0.1307 | 0.1633 | 0.21909 | 0.056362 | 0.098319 | 0.088034 | 0.10954
| Rel.Std.Dev. |  | 4.651% | 3.055% | 3.524% | 5.477% | 1.279% | 2.468% | 2.099% | 2.548%
| PDM3 |  | 1.226% | 0.3295% | 8.652% | -6.2% | 3.376% | -6.591% | -1.627% | 0.8351%
| Z-Score (Absolute) |  | 0.17 | 0.07 | 2.18 | 1.83 | 0.75 | 1.93 | 0.6 | 0.07
| Analytical results for Ge in OREAS 263b (Certified Value 0.12 ppm) | [21 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.8 | 0.24372 | 0.2696 | 0.1 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.20973 | NA | 1.5 | 0.12 | 1.4 | 1.7 | [5 empty] | 1
| 1 | 2 | 1.8 | 0.21834 | 0.2418 | 0.2 | 0.11 | 0.11 | 0.13 | 0.13 | 0.12 | 0.25452 | NA | 1.5 | 0.08 | 1.3 | 1.71 | [5 empty] | 7
| 1 | 3 | 1.8 | 0.24587 | 0.2876 | 0.2 | 0.11 | 0.12 | 0.14 | 0.12 | 0.14 | 0.24626 | NA | 1.5 | 0.1 | 1.3 | 1.63 | [5 empty] | 16
| 1 | 4 | 1.8 | 0.22889 | 0.2842 | 0.3 | 0.11 | 0.11 | 0.13 | 0.12 | 0.07 | 0.20494 | NA | 1.5 | 0.06 | 1.3 | 1.72 | [5 empty] | 0.11739
| 1 | 5 | 1.8 | 0.20198 | 0.1962 | 0.3 | 0.11 | 0.1 | 0.13 | 0.12 | 0.11 | 0.19506 | NA | 1.5 | 0.12 | 1.5 | 1.83 | [5 empty] | 34
| 1 | 6 | 1.7 | 0.21711 | 0.2312 | 0.1 | 0.12 | 0.11 | 0.13 | 0.12 | 0.12 | 0.25844 | NA | 1.5 | 0.1 | 1.4 | 1.75
| Mean |  | 1.7833 | 0.22599 | 0.25177 | 0.2 | 0.11333 | 0.11333 | 0.13 | 0.12167 | 0.11333 | 0.22816 |  | 1.5 | 0.096667 | 1.3667 | 1.7233
| Median |  | 1.8 | 0.22362 | 0.2557 | 0.2 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.22799 |  | 1.5 | 0.1 | 1.35 | 1.715
| Std Dev. |  | 0.040825 | 0.016923 | 0.035395 | 0.089443 | 0.005164 | 0.010328 | 0.0063246 | 0.0040825 | 0.023381 | 0.027979 |  | 0 | 0.023381 | 0.08165 | 0.065625
| Rel.Std.Dev. |  | 2.289% | 7.488% | 14.06% | 44.72% | 4.556% | 9.113% | 4.865% | 3.355% | 20.63% | 12.26% |  | 0% | 24.19% | 5.974% | 3.808%
| PDM3 |  | 1419% | 92.51% | 114.5% | 70.37% | -3.455% | -3.455% | 10.74% | 3.644% | -3.455% | 94.36% |  | 1178% | -17.65% | 1064% | 1368%
| Z-Score (Absolute) |  | NA | 3.24 | 4.08 | NA | 0.4 | 0.4 | 0.13 | 0.13 | 0.4 | 3.31 | NA | NA | 0.94 | NA | 51.71
| Analytical results for Hf in OREAS 263b (Certified Value 3.37 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 3.3 | 3.5054 | 0.6 | 3.6 | 3.5 | 3.5 | 3.6 | 3.5 | 3.3417 | 3.04 | 3.1 | 3.27 | 3.5 | 2.91 | 3.31 | 4 |  |  |  |  | 1
| 1 | 2 | 3.18 | 3.3502 | 2.9 | 3.6 | 3.5 | 3.4 | 3.7 | 3.5 | 3.2486 | 3.01 | 3 | 3.03 | 3.52 | 2.89 | 3.5 | 4 |  |  |  |  | 22
| 1 | 3 | 3.62 | 3.3575 | 2.4 | 3.6 | 3.5 | 3.4 | 3.5 | 3.6 | 3.0877 | 2.91 | 3.1 | 3.17 | 3.53 | 2.84 | 3.37 | 3.8 |  |  |  |  | 16
| 1 | 4 | 3.2 | 3.4333 | 3.1 | 3.6 | 3.5 | 3.3 | 3.5 | 3.6 | 3.3602 | 2.96 | 3 | 3.36 | 3.51 | 2.9 | 3.41 | 3.8 |  |  |  |  | 3.3742
| 1 | 5 | 3.44 | 3.4538 | 3.4 | 3.6 | 3.2 | 3.3 | 3.5 | 3.4 | 3.0943 | 2.89 | 3 | 3.53 | 3.5 | 2.99 | 3.55 | 4 |  |  |  |  | 35
| 1 | 6 | 3.4 | 3.4616 | 0.5 | 3.6 | 3.5 | 3.4 | 3.7 | 3.5 | 3.3602 | 2.94 | 3.1 | 3.38 | 3.75 | 2.89 | 3.49 | 4
| Mean |  | 3.3567 | 3.4269 | 2.15 | 3.6 | 3.45 | 3.3833 | 3.5833 | 3.5167 | 3.2488 | 2.9583 | 3.05 | 3.29 | 3.5517 | 2.9033 | 3.4383 | 3.9333
| Median |  | 3.35 | 3.4435 | 2.65 | 3.6 | 3.5 | 3.4 | 3.55 | 3.5 | 3.2952 | 2.95 | 3.05 | 3.315 | 3.515 | 2.895 | 3.45 | 4
| Std Dev. |  | 0.16561 | 0.061363 | 1.2818 | 0 | 0.12247 | 0.075277 | 0.098319 | 0.075277 | 0.12904 | 0.057764 | 0.054772 | 0.17493 | 0.09786 | 0.048854 | 0.090425 | 0.10328
| Rel.Std.Dev. |  | 4.934% | 1.791% | 59.62% | 0% | 3.55% | 2.225% | 2.744% | 2.141% | 3.972% | 1.953% | 1.796% | 5.317% | 2.755% | 1.683% | 2.63% | 2.626%
| PDM3 |  | -0.5192% | 1.564% | -36.28% | 6.692% | 2.247% | 0.2711% | 6.198% | 4.223% | -3.716% | -12.32% | -9.608% | -2.495% | 5.26% | -13.95% | 1.901% | 16.57%
| Z-Score (Absolute) |  | 0.22 | 0.1 | 5.59 | 0.87 | 0.2 | 0.1 | 0.79 | 0.5 | 0.7 | 1.99 | 1.58 | 0.51 | 0.65 | 2.23 | 0.15 | 2.35
| Analytical results for Hg in OREAS 263b (Indicative Value 0.29 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.15 | 1.21 | [17 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [17 empty] | 2
2
| 1 | 1 | < 1 | 0.3 | 0.26 | [17 empty] | 1
| 1 | 2 | < 1 | 0.29 | 0.32 | [17 empty] | 2
| 1 | 3 | < 1 | 0.29 | 0.3 | [17 empty] | 16
| 1 | 4 | < 1 | 0.29 | 0.28 | [17 empty] | 0.28583
| 1 | 5 | < 1 | 0.25 | 0.29 | [17 empty] | 8
| 1 | 6 | < 1 | 0.28 | 0.28
| Mean |  |  | 0.28333 | 0.28833
| Median |  |  | 0.29 | 0.285
| Std Dev. |  |  | 0.017512 | 0.020412
| Rel.Std.Dev. |  |  | 6.181% | 7.079%
| PDM3 |  |  | -0.8746% | 0.8746%
| Z-Score (Absolute) |  | 28.55 | 0.67 | 0
| Analytical results for Ho in OREAS 263b (Certified Value 0.58 ppm) | [21 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.56 | 0.57002 | 0.6029 | 0.7 | 0.5 | 0.58 | 0.5 | 0.54 | 0.66 | [11 empty] | 1
| 1 | 2 | 0.54 | 0.55824 | 0.60974 | 0.7 | 0.4 | 0.56 | 0.5 | 0.58 | 0.66 | [11 empty] | 23
| 1 | 3 | 0.57 | 0.5299 | 0.58292 | 0.7 | 0.5 | 0.55 | 0.5 | 0.55 | 0.66 | [11 empty] | 16
| 1 | 4 | 0.59 | 0.56112 | 0.61586 | 0.7 | 0.4 | 0.55 | 0.5 | 0.55 | 0.66 | [11 empty] | 0.57601
| 1 | 5 | 0.58 | 0.56219 | 0.5939 | 0.7 | 0.5 | 0.58 | 0.5 | 0.58 | 0.66 | [11 empty] | 36
| 1 | 6 | 0.56 | 0.5562 | 0.62 | 0.7 | 0.5 | 0.57 | 0.5 | 0.56 | 0.66
| Mean |  | 0.56667 | 0.55628 | 0.60422 | 0.7 | 0.46667 | 0.565 | 0.5 | 0.56 | 0.66
| Median |  | 0.565 | 0.55968 | 0.60632 | 0.7 | 0.5 | 0.565 | 0.5 | 0.555 | 0.66
| Std Dev. |  | 0.017512 | 0.013761 | 0.013978 | 1.2162e-16 | 0.05164 | 0.013784 | 0 | 0.016733 | 0
| Rel.Std.Dev. |  | 3.09% | 2.474% | 2.313% | 1.737e-14% | 11.07% | 2.44% | 0% | 2.988% | 0%
| PDM3 |  | -1.622% | -3.426% | 4.897% | 21.53% | -18.98% | -1.912% | -13.2% | -2.78% | 14.58%
| Z-Score (Absolute) |  | 0.03 | 0.15 | 0.67 | 2.32 | 1.69 | 0 | 1.12 | 0.09 | 1.63
| Analytical results for In in OREAS 263b (Certified Value 0.081 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.08 | < 0.1 | 0.11 | 0.082 | 0.08 | 0.086 | 0.077 | 0.083 | 0.080035 | 0.112 | 0.07 | 0.1 | 0.09 | 0.08 | 0.065 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.08 | < 0.1 | 0.14 | 0.078 | 0.075 | 0.086 | 0.073 | 0.091 | 0.07998 | 0.114 | 0.08 | 0.11 | 0.08 | 0.08 | 0.087 | 0.05 |  |  |  |  | 4
| 1 | 3 | 0.07 | < 0.1 | 0.13 | 0.085 | 0.082 | 0.085 | 0.076 | 0.085 | 0.076886 | 0.112 | 0.08 | 0.1 | 0.09 | 0.09 | 0.084 | 0.05 |  |  |  |  | 16
| 1 | 4 | 0.09 | < 0.1 | 0.13 | 0.085 | 0.08 | 0.085 | 0.077 | 0.072 | 0.08184 | 0.114 | 0.07 | 0.1 | 0.09 | 0.08 | 0.062 | 0.05 |  |  |  |  | 0.080845
| 1 | 5 | 0.09 | < 0.1 | 0.14 | 0.086 | 0.08 | 0.086 | 0.07 | 0.077 | 0.077084 | 0.11 | 0.08 | 0.12 | 0.1 | 0.08 | 0.072 | 0.05 |  |  |  |  | 37
| 1 | 6 | 0.07 | < 0.1 | 0.13 | 0.079 | 0.087 | 0.08 | 0.075 | 0.079 | 0.08335 | 0.119 | 0.08 | 0.1 | 0.09 | 0.08 | 0.089 | 0.05
| Mean |  | 0.08 |  | 0.13 | 0.0825 | 0.080667 | 0.084667 | 0.074667 | 0.081167 | 0.079863 | 0.1135 | 0.076667 | 0.105 | 0.09 | 0.081667 | 0.0765 | 0.05
| Median |  | 0.08 |  | 0.13 | 0.0835 | 0.08 | 0.0855 | 0.0755 | 0.081 | 0.080008 | 0.113 | 0.08 | 0.1 | 0.09 | 0.08 | 0.078 | 0.05
| Std Dev. |  | 0.0089443 |  | 0.010954 | 0.0033912 | 0.0038816 | 0.0023381 | 0.0027325 | 0.0066458 | 0.0025572 | 0.0030822 | 0.005164 | 0.0083666 | 0.0063246 | 0.0040825 | 0.011709 | 7.6012e-18
| Rel.Std.Dev. |  | 11.18% |  | 8.427% | 4.111% | 4.812% | 2.762% | 3.66% | 8.188% | 3.202% | 2.716% | 6.736% | 7.968% | 7.027% | 4.999% | 15.31% | 1.52e-14%
| PDM3 |  | -1.045% |  | 60.8% | 2.047% | -0.2207% | 4.727% | -7.642% | 0.3978% | -1.215% | 40.39% | -5.168% | 29.88% | 11.32% | 1.016% | -5.375% | -38.15%
| Z-Score (Absolute) |  | 0.14 | 4.81 | 7.64 | 0.25 | 0.04 | 0.58 | 0.97 | 0.04 | 0.16 | 5.07 | 0.66 | 3.75 | 1.41 | 0.12 | 0.69 | 4.81
| Analytical results for K in OREAS 263b (Certified Value 2.03 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.0026 | 2.0078 | 1.9106 | 2.24 | 2.14 | 2.01 | 2.07 | 1.98 | 2.01 | 1.96 | 2.0912 | 2.0137 | 2.1 | 1.75 | 2.1945 | 2.1 | 2.14 | 1.98 | 2 | 1.8815 | 1
| 1 | 2 | 2.0363 | 2.0458 | 1.9136 | 2.38 | 2.15 | 2 | 2.11 | 1.98 | 1.99 | 2.03 | 2.0367 | 2.0112 | 2.1 | 2.08 | 2.2195 | 2.09 | 2.08 | 1.95 | 2.01 | 1.9451 | #N/A
| 1 | 3 | 2.0194 | 2.0406 | 1.9211 | 2.2 | 2.13 | 2.01 | 2.07 | 2 | 2 | 2.02 | 2.0484 | 1.9771 | 2.01 | 1.9 | 2.1533 | 2.11 | 1.96 | 1.96 | 2.05 | 2.0654 | 16
| 1 | 4 | 2.0093 | 1.9898 | 1.918 | 2.33 | 2.15 | 2.03 | 2.1 | 2.01 | 2.01 | 2.08 | 2.0315 | 1.9915 | 2.09 | 2.13 | 2.1979 | 2.08 | 2.08 | 1.97 | 2.02 | 2.0849 | 2.0324
| 1 | 5 | 2.021 | 2.0456 | 1.9226 | 2.3 | 2.1 | 1.99 | 2.12 | 1.99 | 2.04 | 2.04 | 2.0819 | 2.0206 | 2.04 | 1.89 | 2.2139 | 2.13 | 2.23 | 1.93 | 2.02 | 2.0581 | 38
| 1 | 6 | 2.027 | 2.0614 | 1.9287 | 2.36 | 2.13 | 2.04 | 2.08 | 1.98 | 2.03 | 2.03 | 2.0542 | 1.9959 | 2.08 | 2.08 | 2.1849 | 2.13 | 2.04 | 1.96 | 2.03 | 2.1077
| Mean |  | 2.0193 | 2.0319 | 1.9191 | 2.3017 | 2.1333 | 2.0133 | 2.0917 | 1.99 | 2.0133 | 2.0267 | 2.0573 | 2.0017 | 2.07 | 1.9717 | 2.194 | 2.1067 | 2.0883 | 1.9583 | 2.0217 | 2.0238
| Median |  | 2.0202 | 2.0431 | 1.9195 | 2.315 | 2.135 | 2.01 | 2.09 | 1.985 | 2.01 | 2.03 | 2.0513 | 2.0036 | 2.085 | 1.99 | 2.1962 | 2.105 | 2.08 | 1.96 | 2.02 | 2.0617
| Std Dev. |  | 0.012084 | 0.027132 | 0.0065309 | 0.069976 | 0.018619 | 0.018619 | 0.02137 | 0.012649 | 0.018619 | 0.038816 | 0.024213 | 0.016336 | 0.036878 | 0.14798 | 0.023669 | 0.020656 | 0.091305 | 0.017224 | 0.017224 | 0.089582
| Rel.Std.Dev. |  | 0.5984% | 1.335% | 0.3403% | 3.04% | 0.8728% | 0.9248% | 1.022% | 0.6356% | 0.9248% | 1.915% | 1.177% | 0.8161% | 1.782% | 7.505% | 1.079% | 0.9805% | 4.372% | 0.8795% | 0.852% | 4.426%
| PDM3 |  | -0.6438% | -0.02454% | -5.573% | 13.25% | 4.969% | -0.9357% | 2.919% | -2.084% | -0.9357% | -0.2797% | 1.229% | -1.51% | 1.853% | -2.986% | 7.954% | 3.657% | 2.755% | -3.642% | -0.5257% | -0.422%
| Z-Score (Absolute) |  | 0.1 | 0.11 | 1.79 | 4.66 | 1.82 | 0.2 | 1.12 | 0.59 | 0.2 | 0.02 | 0.54 | 0.4 | 0.75 | 0.9 | 2.85 | 1.37 | 1.06 | 1.13 | 0.06 | 0.02
| Analytical results for La in OREAS 263b (Certified Value 40.3 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  |  | 1
2
| 1 | 1 | 41.37 | 38.3803 | 45.037 | 44.6 | 36.8 | 42.4 | 40.9 | 40.1 | 40.4 | 40.1 | 31.8 | 40.9 | 42.62 | 39.9 | 36.8 | 38.64 | 38.5 | 37.532 |  |  | 1
| 1 | 2 | 41.22 | 38.205 | 46.2364 | 43.7 | 37.1 | 42.3 | 41.7 | 39.4 | 39.1 | 42.3 | 31.5 | 39.6 | 42.79 | 39.2 | 37.8 | 38.07 | 39.3 | 37.959 |  |  | 25
| 1 | 3 | 41.15 | 38.2423 | 46.8461 | 44.5 | 37.6 | 43.1 | 40.1 | 39.6 | 37.8 | 44.6 | 30.7 | 35.9 | 41.97 | 38.2 | 36.9 | 38.47 | 38.9 | 35.233 |  |  | 16
| 1 | 4 | 41.51 | 39.3915 | 46.5431 | 43.1 | 36.6 | 43.2 | 40.3 | 41.1 | 39.2 | 42.5 | 31.3 | 43.6 | 42.38 | 40.4 | 36.3 | 38.41 | 39 | 35.76 |  |  | 40.2916
| 1 | 5 | 41.66 | 38.8001 | 46.342 | 43.2 | 36.4 | 43.8 | 37.3 | 39.5 | 39.2 | 41.8 | 30.3 | 39 | 42.66 | 39.8 | 37.1 | 37.8 | 39.7 | 38.307 |  |  | 39
| 1 | 6 | 41.12 | 38.4795 | 46.5549 | 42.8 | 35.4 | 42.3 | 41.1 | 41 | 39.6 | 45.1 | 31.2 | 35.1 | 43.15 | 39.3 | 36.5 | 37.76 | 39.7 | 38.059
| Mean |  | 41.3383 | 38.5831 | 46.2599 | 43.65 | 36.65 | 42.85 | 40.2333 | 40.1167 | 39.2167 | 42.7333 | 31.1333 | 39.0167 | 42.595 | 39.4667 | 36.9 | 38.1917 | 39.1833 | 37.1417
| Median |  | 41.295 | 38.4299 | 46.4425 | 43.45 | 36.7 | 42.75 | 40.6 | 39.85 | 39.2 | 42.4 | 31.25 | 39.3 | 42.64 | 39.55 | 36.85 | 38.24 | 39.15 | 37.7455
| Std Dev. |  | 0.21498 | 0.44978 | 0.6346 | 0.75565 | 0.74229 | 0.61563 | 1.5475 | 0.76267 | 0.84479 | 1.8511 | 0.5465 | 3.1606 | 0.39693 | 0.75807 | 0.52536 | 0.36897 | 0.47504 | 1.3093
| Rel.Std.Dev. |  | 0.5201% | 1.166% | 1.372% | 1.731% | 2.025% | 1.437% | 3.846% | 1.901% | 2.154% | 4.332% | 1.755% | 8.101% | 0.9319% | 1.921% | 1.424% | 0.9661% | 1.212% | 3.525%
| PDM3 |  | 2.598% | -4.24% | 14.81% | 8.335% | -9.038% | 6.35% | -0.1447% | -0.4342% | -2.668% | 6.06% | -22.73% | -3.164% | 5.717% | -2.047% | -8.418% | -5.212% | -2.751% | -7.818%
| Z-Score (Absolute) |  | 0.64 | 0.24 | 2.22 | 1.38 | 0.86 | 1.13 | 0.29 | 0.25 | 0.04 | 1.09 | 2.64 | 0.1 | 1.05 | 0.04 | 0.78 | 0.37 | 0.05 | 0.71
| Analytical results for Li in OREAS 263b (Certified Value 37.9 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 38.5 | 33.0742 | 32.5361 | 38 | 37.3 | 37.2 | 38.4 | 38.2 | 37.9 | 36.7 | 38.5106 | 34 | 38.1 | 35.3 | 39.1 | 37.1 | 34.4 | 41.51 | 40 | 34.223 | 1
| 1 | 2 | 37.4 | 33.3792 | 32.6397 | 39 | 37.5 | 37.2 | 39.2 | 38.1 | 37.5 | 38.6 | 38.6366 | 34 | 38.1 | 34.5 | 38.6 | 37.3 | 36.1 | 39.85 | 30 | 34.64 | #N/A
| 1 | 3 | 38.6 | 33.2342 | 32.5292 | 39 | 36.9 | 37.3 | 38.7 | 37.9 | 37.7 | 38.2 | 39.289 | 34 | 38.1 | 39.6 | 38.9 | 37.5 | 35.5 | 41.83 | 40 | 34.428 | 16
| 1 | 4 | 38.5 | 32.9914 | 32.471 | 39 | 37.1 | 37.7 | 39.1 | 38.4 | 38 | 37.9 | 38.9361 | 34 | 38.1 | 37.2 | 39.5 | 36.8 | 39.4 | 41.17 | 40 | 34.935 | 37.9046
| 1 | 5 | 38.1 | 33.4077 | 32.6228 | 39 | 37 | 37.2 | 39.5 | 38.3 | 40.1 | 38.8 | 39.1161 | 35 | 38.1 | 39.9 | 39.3 | 37.9 | 36.4 | 40.72 | 30 | 34.921 | 40
| 1 | 6 | 37.6 | 33.181 | 32.5671 | 39 | 37.4 | 37.9 | 38.7 | 38.1 | 38.2 | 37.7 | 39.536 | 35 | 38.1 | 38.9 | 39 | 37.6 | 38.6 | 40.36 | 40 | 35.021
| Mean |  | 38.1167 | 33.2113 | 32.561 | 38.8333 | 37.2 | 37.4167 | 38.9333 | 38.1667 | 38.2333 | 37.9833 | 39.0041 | 34.3333 | 38.1 | 37.5667 | 39.0667 | 37.3667 | 36.7333 | 40.9067 | 36.6667 | 34.6947
| Median |  | 38.3 | 33.2076 | 32.5516 | 39 | 37.2 | 37.25 | 38.9 | 38.15 | 37.95 | 38.05 | 39.0261 | 34 | 38.1 | 38.05 | 39.05 | 37.4 | 36.25 | 40.945 | 40 | 34.7805
| Std Dev. |  | 0.51153 | 0.16454 | 0.062914 | 0.40825 | 0.23664 | 0.30605 | 0.40332 | 0.17512 | 0.94587 | 0.75211 | 0.38981 | 0.5164 | 0 | 2.2818 | 0.31411 | 0.38816 | 1.9012 | 0.73975 | 5.164 | 0.32002
| Rel.Std.Dev. |  | 1.342% | 0.4954% | 0.1932% | 1.051% | 0.6361% | 0.818% | 1.036% | 0.4588% | 2.474% | 1.98% | 0.9994% | 1.504% | 0% | 6.074% | 0.804% | 1.039% | 5.176% | 1.808% | 14.08% | 0.9224%
| PDM3 |  | 0.5596% | -12.38% | -14.1% | 2.45% | -1.859% | -1.287% | 2.714% | 0.6915% | 0.8674% | 0.2078% | 2.901% | -9.422% | 0.5156% | -0.8914% | 3.066% | -1.419% | -3.09% | 7.92% | -3.266% | -8.468%
| Z-Score (Absolute) |  | 0.22 | 2.93 | 3.35 | 0.68 | 0.37 | 0.23 | 0.74 | 0.25 | 0.29 | 0.13 | 0.79 | 2.21 | 0.21 | 0.13 | 0.83 | 0.26 | 0.67 | 2.01 | 0.71 | 1.98
| Analytical results for Lu in OREAS 263b (Certified Value 0.27 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.25 | 0.26556 | 0.25766 | 0.3 | 0.2 | 0.3 | 0.2 | 0.24 | 0.28 | [11 empty] | 1
| 1 | 2 | 0.29 | 0.23956 | 0.24803 | 0.3 | 0.2 | 0.3 | 0.2 | 0.26 | 0.28 | [11 empty] | 5
| 1 | 3 | 0.32 | 0.24797 | 0.29159 | 0.3 | 0.2 | 0.3 | 0.2 | 0.23 | 0.28 | [11 empty] | 16
| 1 | 4 | 0.25 | 0.26447 | 0.27224 | 0.3 | 0.2 | 0.29 | 0.2 | 0.22 | 0.3 | [11 empty] | 0.2664
| 1 | 5 | 0.25 | 0.24228 | 0.29024 | 0.3 | 0.2 | 0.28 | 0.2 | 0.23 | 0.28 | [11 empty] | 41
| 1 | 6 | 0.25 | 0.2558 | 0.27701 | 0.3 | 0.2 | 0.28 | 0.2 | 0.24 | 0.3
| Mean |  | 0.26833 | 0.25261 | 0.2728 | 0.3 | 0.2 | 0.29167 | 0.2 | 0.23667 | 0.28667
| Median |  | 0.25 | 0.25188 | 0.27463 | 0.3 | 0.2 | 0.295 | 0.2 | 0.235 | 0.28
| Std Dev. |  | 0.029944 | 0.01111 | 0.017424 | 0 | 3.0405e-17 | 0.0098319 | 3.0405e-17 | 0.013663 | 0.010328
| Rel.Std.Dev. |  | 11.16% | 4.398% | 6.387% | 0% | 1.52e-14% | 3.371% | 1.52e-14% | 5.773% | 3.603%
| PDM3 |  | 0.7257% | -5.178% | 2.4% | 12.61% | -24.92% | 9.484% | -24.92% | -11.16% | 7.608%
| Z-Score (Absolute) |  | 0.09 | 0.71 | 0.09 | NA | NA | 0.84 | NA | 1.34 | 0.64
| Analytical results for Mg in OREAS 263b (Certified Value 0.527 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.5213 | 0.51395 | 0.55685 | 0.48 | 0.52 | 0.51 | 0.52 | 0.53 | 0.49 | 0.47 | 0.56259 | 0.5102 | 0.491 | 0.53 | 0.5345 | 0.51 | 0.508 | 0.58 | 0.6 | 0.55019 | 1
| 1 | 2 | 0.5369 | 0.51462 | 0.55555 | 0.49 | 0.52 | 0.51 | 0.52 | 0.53 | 0.48 | 0.48 | 0.53941 | 0.5101 | 0.503 | 0.55 | 0.5484 | 0.51 | 0.498 | 0.57 | 0.61 | 0.55418 | #N/A
| 1 | 3 | 0.5323 | 0.5089 | 0.55561 | 0.49 | 0.52 | 0.52 | 0.52 | 0.53 | 0.48 | 0.47 | 0.54255 | 0.5034 | 0.509 | 0.53 | 0.5461 | 0.51 | 0.461 | 0.58 | 0.61 | 0.54831 | 16
| 1 | 4 | 0.5256 | 0.51391 | 0.55368 | 0.49 | 0.53 | 0.51 | 0.53 | 0.54 | 0.49 | 0.49 | 0.56298 | 0.5087 | 0.5 | 0.58 | 0.5525 | 0.5 | 0.498 | 0.57 | 0.6 | 0.55338 | 0.52699
| 1 | 5 | 0.5436 | 0.51736 | 0.55437 | 0.48 | 0.51 | 0.51 | 0.53 | 0.53 | 0.5 | 0.47 | 0.55315 | 0.5123 | 0.511 | 0.56 | 0.5428 | 0.51 | 0.535 | 0.57 | 0.61 | 0.55446 | 42
| 1 | 6 | 0.5421 | 0.5117 | 0.55376 | 0.49 | 0.52 | 0.52 | 0.52 | 0.53 | 0.49 | 0.49 | 0.54711 | 0.512 | 0.507 | 0.54 | 0.5482 | 0.51 | 0.487 | 0.57 | 0.61 | 0.55119
| Mean |  | 0.53363 | 0.51341 | 0.55497 | 0.48667 | 0.52 | 0.51333 | 0.52333 | 0.53167 | 0.48833 | 0.47833 | 0.5513 | 0.50945 | 0.5035 | 0.54833 | 0.54542 | 0.50833 | 0.49783 | 0.57333 | 0.60667 | 0.55195
| Median |  | 0.5346 | 0.51393 | 0.55496 | 0.49 | 0.52 | 0.51 | 0.52 | 0.53 | 0.49 | 0.475 | 0.55013 | 0.51015 | 0.505 | 0.545 | 0.54715 | 0.51 | 0.498 | 0.57 | 0.61 | 0.55228
| Std Dev. |  | 0.0089455 | 0.0028594 | 0.0012471 | 0.005164 | 0.0063246 | 0.005164 | 0.005164 | 0.0040825 | 0.0075277 | 0.0098319 | 0.010029 | 0.0032488 | 0.0073144 | 0.019408 | 0.0062179 | 0.0040825 | 0.024326 | 0.005164 | 0.005164 | 0.0024592
| Rel.Std.Dev. |  | 1.676% | 0.5569% | 0.2247% | 1.061% | 1.216% | 1.006% | 0.9867% | 0.7679% | 1.542% | 2.055% | 1.819% | 0.6377% | 1.453% | 3.539% | 1.14% | 0.8031% | 4.886% | 0.9007% | 0.8512% | 0.4455%
| PDM3 |  | 1.261% | -2.577% | 5.309% | -7.652% | -1.326% | -2.591% | -0.6938% | 0.8875% | -7.335% | -9.233% | 4.613% | -3.328% | -4.457% | 4.05% | 3.497% | -3.54% | -5.533% | 8.794% | 15.12% | 4.736%
| Z-Score (Absolute) |  | 0.34 | 0.23 | 0.94 | 0.99 | 0.05 | 0.24 | 0.05 | 0.28 | 0.94 | 1.23 | 0.84 | 0.35 | 0.51 | 0.76 | 0.67 | 0.38 | 0.68 | 1.46 | 2.41 | 0.86
| Analytical results for Mn in OREAS 263b (Certified Value 0.017 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.0168 | 0.017271 | 0.017973 | 0.0164 | 0.0168 | 0.0163 | 0.0167 | 0.0176 | 0.0169 | 0.0172 | 0.017189 | 0.0169 | 0.0162 | 0.0184 | 0.0182 | 0.0177 | 0.0169 | 0.01785 | 0.0172 | 0.015885 | 1
| 1 | 2 | 0.0171 | 0.017125 | 0.017997 | 0.0176 | 0.0182 | 0.0165 | 0.017 | 0.0178 | 0.0168 | 0.0172 | 0.017377 | 0.0168 | 0.0163 | 0.0186 | 0.0184 | 0.0177 | 0.0165 | 0.01772 | 0.0174 | 0.01392 | 27
| 1 | 3 | 0.0168 | 0.017117 | 0.018694 | 0.0163 | 0.0178 | 0.0168 | 0.0169 | 0.0177 | 0.0167 | 0.017 | 0.017079 | 0.0167 | 0.0166 | 0.0183 | 0.018 | 0.0176 | 0.0153 | 0.01793 | 0.0178 | 0.012543 | 16
| 1 | 4 | 0.0168 | 0.017136 | 0.018256 | 0.0165 | 0.0172 | 0.0167 | 0.017 | 0.018 | 0.0169 | 0.0179 | 0.018157 | 0.0168 | 0.0163 | 0.0201 | 0.0182 | 0.0173 | 0.0166 | 0.01782 | 0.0174 | 0.0061796 | 0.017356
| 1 | 5 | 0.0169 | 0.017218 | 0.018577 | 0.0172 | 0.0174 | 0.0165 | 0.0173 | 0.0178 | 0.0173 | 0.0169 | 0.017746 | 0.0168 | 0.016 | 0.0191 | 0.0184 | 0.0179 | 0.018 | 0.01775 | 0.018 | 0.015453 | 43
| 1 | 6 | 0.0171 | 0.017211 | 0.018518 | 0.0186 | 0.0167 | 0.017 | 0.0169 | 0.0176 | 0.0169 | 0.0178 | 0.018384 | 0.0169 | 0.0166 | 0.0185 | 0.0183 | 0.0177 | 0.0162 | 0.01765 | 0.018 | 0.016707
| Mean |  | 0.016917 | 0.01718 | 0.018336 | 0.0171 | 0.01735 | 0.016633 | 0.016967 | 0.01775 | 0.016917 | 0.017333 | 0.017655 | 0.016817 | 0.016333 | 0.018833 | 0.01825 | 0.01765 | 0.016583 | 0.017787 | 0.017633 | 0.013448
| Median |  | 0.01685 | 0.017174 | 0.018387 | 0.01685 | 0.0173 | 0.0166 | 0.01695 | 0.01775 | 0.0169 | 0.0172 | 0.017562 | 0.0168 | 0.0163 | 0.01855 | 0.01825 | 0.0177 | 0.01655 | 0.017785 | 0.0176 | 0.014686
| Std Dev. |  | 0.0001472 | 6.263e-05 | 0.00030718 | 0.00089443 | 0.00057879 | 0.00025033 | 0.00019664 | 0.00015166 | 0.00020412 | 0.00041793 | 0.00053255 | 7.5277e-05 | 0.00023381 | 0.0006802 | 0.00015166 | 0.00019748 | 0.00088412 | 0.00010013 | 0.00034448 | 0.0038588
| Rel.Std.Dev. |  | 0.8701% | 0.3646% | 1.675% | 5.231% | 3.336% | 1.505% | 1.159% | 0.8544% | 1.207% | 2.411% | 3.016% | 0.4476% | 1.431% | 3.612% | 0.831% | 1.119% | 5.331% | 0.563% | 1.954% | 28.69%
| PDM3 |  | -2.533% | -1.017% | 5.643% | -1.477% | -0.03671% | -4.166% | -2.245% | 2.268% | -2.533% | -0.1327% | 1.723% | -3.11% | -5.894% | 8.51% | 5.149% | 1.692% | -4.454% | 2.479% | 1.596% | -22.52%
| Z-Score (Absolute) |  | 0.55 | 0.12 | 1.74 | 0.25 | 0.15 | 1 | 0.47 | 0.79 | 0.55 | 0.12 | 0.64 | 0.71 | 1.48 | 2.54 | 1.6 | 0.63 | 1.08 | 0.85 | 0.61 | 6.12
| Analytical results for Mo in OREAS 263b (Certified Value 2.36 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 2.5 | 2.1903 | < 1 | 2.3 | 2.39 | 2.5 | 2.41 | 2.43 | 2.28 | 1.9868 | 2.2 | 2.4 | 1.97 | 2.5 | 2.35 | 1.6 | 1.94 | 2.5 | 2.681 |  | 1
| 1 | 2 | 2.5 | 2.1182 | < 1 | 2.2 | 2.38 | 2.49 | 2.47 | 2.37 | 2.31 | 2.072 | 2.2 | 2.4 | 2.79 | 2.4 | 2.28 | 1.8 | 1.91 | 2.5 | 2.649 |  | 28
| 1 | 3 | 2.4 | 2.0522 | < 1 | 2.2 | 2.47 | 2.25 | 2.52 | 2.32 | 2.6 | 2.1448 | 2.3 | 2.4 | 2.31 | 2.4 | 2.32 | 1.7 | 1.93 | 2.5 | 2.809 |  | 16
| 1 | 4 | 2.3 | 2.1089 | < 1 | 2.1 | 2.96 | 2.3 | 2.55 | 2.42 | 3.11 | 2.1718 | 2.2 | 2.4 | 2.05 | 2.5 | 2.26 | 1.6 | 1.9 | 2.5 | 2.559 |  | 2.3601
| 1 | 5 | 2.4 | 2.1506 | < 1 | 2.4 | 2.3 | 2.18 | 2.44 | 2.33 | 2.43 | 2.0351 | 2.2 | 2.3 | 2.64 | 2.5 | 2.33 | 2.1 | 1.93 | 2.5 | 2.924 |  | 44
| 1 | 6 | 2.4 | 2.0904 | < 1 | 2.1 | 2.42 | 2.36 | 2.58 | 2.45 | 2.45 | 2.1163 | 2.3 | 2.3 | 2.47 | 2.4 | 2.27 | 2 | 1.56 | 2.5 | 2.621
| Mean |  | 2.4167 | 2.1184 |  | 2.2167 | 2.4867 | 2.3467 | 2.495 | 2.3867 | 2.53 | 2.0878 | 2.2333 | 2.3667 | 2.3717 | 2.45 | 2.3017 | 1.8 | 1.8617 | 2.5 | 2.7072
| Median |  | 2.4 | 2.1136 |  | 2.2 | 2.405 | 2.33 | 2.495 | 2.395 | 2.44 | 2.0941 | 2.2 | 2.4 | 2.39 | 2.45 | 2.3 | 1.75 | 1.92 | 2.5 | 2.665
| Std Dev. |  | 0.075277 | 0.047876 |  | 0.1169 | 0.23847 | 0.12925 | 0.065955 | 0.05465 | 0.30614 | 0.069763 | 0.05164 | 0.05164 | 0.32412 | 0.054772 | 0.03656 | 0.20976 | 0.14851 | 0 | 0.13478
| Rel.Std.Dev. |  | 3.115% | 2.26% |  | 5.274% | 9.59% | 5.508% | 2.643% | 2.29% | 12.1% | 3.341% | 2.312% | 2.182% | 13.67% | 2.236% | 1.588% | 11.65% | 7.978% | 0% | 4.979%
| PDM3 |  | 2.398% | -10.24% |  | -6.076% | 5.364% | -0.5677% | 5.717% | 1.127% | 7.2% | -11.54% | -5.37% | 0.2797% | 0.4916% | 3.811% | -2.474% | -23.73% | -21.12% | 5.929% | 14.71%
| Z-Score (Absolute) |  | 0.25 | 1.26 | 9.44 | 0.76 | 0.61 | 0.1 | 0.65 | 0.1 | 0.83 | 1.41 | 0.67 | 0 | 0.03 | 0.42 | 0.33 | 2.87 | 2.55 | 0.67 | 1.72
| Analytical results for Na in OREAS 263b (Certified Value 0.115 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.1136 | 0.11454 | 0.13199 | 0.14 | 0.12 | 0.11 | 0.11 | 0.12 | 0.12 | 0.19 | 0.11338 | 0.083 | 0.108 | 0.151 | 0.1223 | 0.109 | 0.116 | 0.16 | 0.12 | 0.10465 | 1
| 1 | 2 | 0.1152 | 0.1138 | 0.13224 | 0.14 | 0.12 | 0.11 | 0.11 | 0.12 | 0.12 | 0.21 | 0.10972 | 0.0825 | 0.101 | 0.152 | 0.1234 | 0.108 | 0.0963 | 0.17 | 0.12 | 0.10711 | #N/A
| 1 | 3 | 0.1156 | 0.11178 | 0.13319 | 0.15 | 0.12 | 0.12 | 0.11 | 0.12 | 0.12 | 0.21 | 0.11023 | 0.0813 | 0.105 | 0.159 | 0.125 | 0.11 | 0.0903 | 0.16 | 0.11 | 0.10869 | 16
| 1 | 4 | 0.1137 | 0.11269 | 0.13283 | 0.15 | 0.12 | 0.12 | 0.11 | 0.12 | 0.12 | 0.2 | 0.11343 | 0.0827 | 0.102 | 0.16 | 0.1245 | 0.109 | 0.112 | 0.19 | 0.11 | 0.10006 | 0.11487
| 1 | 5 | 0.1181 | 0.11522 | 0.13451 | 0.14 | 0.12 | 0.11 | 0.11 | 0.12 | 0.12 | 0.21 | 0.11037 | 0.0825 | 0.103 | 0.166 | 0.1218 | 0.113 | 0.115 | 0.16 | 0.12 | 0.10796 | 45
| 1 | 6 | 0.117 | 0.11562 | 0.13301 | 0.14 | 0.12 | 0.12 | 0.11 | 0.12 | 0.12 | 0.2 | 0.11295 | 0.0833 | 0.102 | 0.165 | 0.1248 | 0.11 | 0.1 | 0.18 | 0.11 | 0.10671
| Mean |  | 0.11553 | 0.11394 | 0.13296 | 0.14333 | 0.12 | 0.115 | 0.11 | 0.12 | 0.12 | 0.20333 | 0.11168 | 0.08255 | 0.1035 | 0.15883 | 0.12363 | 0.10983 | 0.10493 | 0.17 | 0.115 | 0.10586
| Median |  | 0.1154 | 0.11417 | 0.13292 | 0.14 | 0.12 | 0.115 | 0.11 | 0.12 | 0.12 | 0.205 | 0.11166 | 0.0826 | 0.1025 | 0.1595 | 0.12395 | 0.1095 | 0.106 | 0.165 | 0.115 | 0.10691
| Std Dev. |  | 0.0017862 | 0.0014884 | 0.00088758 | 0.005164 | 0 | 0.0054772 | 0 | 0 | 0 | 0.008165 | 0.0017453 | 0.00068629 | 0.0025884 | 0.0063061 | 0.0013545 | 0.0017224 | 0.010833 | 0.012649 | 0.0054772 | 0.0031586
| Rel.Std.Dev. |  | 1.546% | 1.306% | 0.6675% | 3.603% | 0% | 4.763% | 0% | 0% | 0% | 4.016% | 1.563% | 0.8314% | 2.501% | 3.97% | 1.096% | 1.568% | 10.32% | 7.441% | 4.763% | 2.984%
| PDM3 |  | 0.5778% | -0.8079% | 15.75% | 24.78% | 4.466% | 0.1135% | -4.239% | 4.466% | 4.466% | 77.01% | -2.776% | -28.14% | -9.898% | 38.27% | 7.629% | -4.384% | -8.65% | 47.99% | 0.1135% | -7.84%
| Z-Score (Absolute) |  | 0.03 | 0.13 | 1.73 | 2.74 | 0.46 | 0.03 | 0.51 | 0.46 | 0.46 | 8.61 | 0.35 | 3.2 | 1.15 | 4.26 | 0.82 | 0.53 | 1.01 | 5.35 | 0.03 | 0.92
| Analytical results for Nb in OREAS 263b (Certified Value 10.1 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
1
| 1 | 1 | 9.71 | 10.0595 | 9.3 | 7 | 11.8 | 10.5 | 11 | 11.3 | 10.2 | 7.2187 | NA | 6.4 | 9.65 | 9.82 | 8.4 | 11.1 | 12 |  |  |  | 1
| 1 | 2 | 10.18 | 9.983 | 8.1 | 7 | 11.6 | 10.5 | 11 | 11.1 | 10.4 | 7.775 | NA | 6.4 | 8.88 | 10.21 | 8.54 | 11.5 | 12 |  |  |  | 29
| 1 | 3 | 9.96 | 9.7698 | 8.9 | 7 | 11.8 | 10.2 | 11 | 10.6 | 11 | 7.9357 | NA | 6.4 | 9.42 | 10.57 | 8.53 | 10.4 | 12 |  |  |  | 16
| 1 | 4 | 10.03 | 9.7973 | 9 | 7 | 11.8 | 10.5 | 11.3 | 11.1 | 10.6 | 7.1552 | NA | 6 | 10.22 | 10.45 | 8.35 | 11.2 | 12 |  |  |  | 10.0686
| 1 | 5 | 9.98 | 10.1302 | 8.9 | 7 | 11.9 | 9.5 | 11 | 10.8 | 10.4 | 6.7892 | NA | 6 | 10.8 | 9.94 | 8.64 | 11.4 | 12 |  |  |  | 46
| 1 | 6 | 10.17 | 10.4128 | 9.3 | 6 | 11.7 | 10.7 | 11.2 | 11.1 | 10.4 | 6.8658 | NA | 6.2 | 10.13 | 10.64 | 8.56 | 11.3 | 12
| Mean |  | 10.005 | 10.0254 | 8.9167 | 6.8333 | 11.7667 | 10.3167 | 11.0833 | 11 | 10.5 | 7.29 |  | 6.2333 | 9.85 | 10.2717 | 8.5033 | 11.15 | 12
| Median |  | 10.005 | 10.0213 | 8.95 | 7 | 11.8 | 10.5 | 11 | 11.1 | 10.4 | 7.187 |  | 6.3 | 9.89 | 10.33 | 8.535 | 11.25 | 12
| Std Dev. |  | 0.17213 | 0.23722 | 0.44008 | 0.40825 | 0.10328 | 0.43089 | 0.13292 | 0.25298 | 0.27568 | 0.47034 |  | 0.19664 | 0.67581 | 0.33891 | 0.10783 | 0.3937 | 0
| Rel.Std.Dev. |  | 1.72% | 2.366% | 4.935% | 5.974% | 0.8777% | 4.177% | 1.199% | 2.3% | 2.626% | 6.452% |  | 3.155% | 6.861% | 3.299% | 1.268% | 3.531% | 0%
| PDM3 |  | -0.6319% | -0.4288% | -11.44% | -32.13% | 16.86% | 2.464% | 10.08% | 9.25% | 4.284% | -27.6% |  | -38.09% | -2.171% | 2.017% | -15.55% | 10.74% | 19.18%
| Z-Score (Absolute) |  | 0.11 | 0.09 | 0.93 | NA | 1.22 | 0.13 | 0.71 | 0.64 | 0.27 | 2.16 | NA | 2.96 | 0.23 | 0.09 | 1.24 | 0.76 | NA
Results from laboratory 1.08 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 263b (Certified Value 33.1 ppm) | [21 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 1
2
| 1 | 1 | 33.06 | 32.2677 | 39.4374 | 33.7 | 28.7 | 35.34 | 31.6 | 32.4 | 35.9 | [11 empty] | 1
| 1 | 2 | 33.1 | 31.3573 | 39.5104 | 34.7 | 26.7 | 35.61 | 31.5 | 33.2 | 35.6 | [11 empty] | 6
| 1 | 3 | 32.86 | 31.6107 | 39.669 | 34.7 | 31.5 | 34.94 | 30.6 | 32.5 | 35.4 | [11 empty] | 16
| 1 | 4 | 32.74 | 32.3623 | 39.6804 | 33.1 | 28.5 | 34.9 | 31.8 | 31.2 | 35.6 | [11 empty] | 33.0779
| 1 | 5 | 32.97 | 32.5248 | 39.5378 | 33.7 | 34.2 | 35.19 | 31.9 | 32.7 | 34.6 | [11 empty] | 47
| 1 | 6 | 32.77 | 32.7045 | 39.4821 | 33.7 | 31.8 | 35.61 | 31.6 | 32 | 35.4
| Mean |  | 32.9167 | 32.1379 | 39.5529 | 33.9333 | 30.2333 | 35.265 | 31.5 | 32.3333 | 35.4167
| Median |  | 32.915 | 32.315 | 39.5241 | 33.7 | 30.1 | 35.265 | 31.6 | 32.45 | 35.5
| Std Dev. |  | 0.15029 | 0.53389 | 0.10011 | 0.6377 | 2.7435 | 0.31246 | 0.46476 | 0.6802 | 0.44008
| Rel.Std.Dev. |  | 0.4566% | 1.661% | 0.2531% | 1.879% | 9.074% | 0.886% | 1.475% | 2.104% | 1.243%
| PDM3 |  | -0.4873% | -2.842% | 19.58% | 2.586% | -8.599% | 6.612% | -4.77% | -2.251% | 7.071%
| Z-Score (Absolute) |  | 0 | 0.37 | 3.16 | 0.48 | 1.28 | 1.12 | 0.67 | 0.28 | 1.19
| Analytical results for Ni in OREAS 263b (Certified Value 44 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 43.8 | 40.2631 | 46.1084 | 48 | 42.2 | 43.5 | 44.5 | 43.7 | 44.9 | 42.8 | 43.6796 | 40.5 | 43 | 41.1 | 44.3 | 44.4 | 40.8 | 43.32 | 54 | 43.877 | 1
| 1 | 2 | 44.8 | 39.6175 | 45.5653 | 48 | 41.1 | 42.4 | 43.9 | 43.7 | 45.6 | 44 | 41.5863 | 38.8 | 43 | 38.6 | 45 | 44.1 | 43.1 | 44.23 | 54 | 44.191 | 14
| 1 | 3 | 45 | 39.4572 | 48.801 | 49 | 40.8 | 44.1 | 41.5 | 43.9 | 44.9 | 46.4 | 41.8407 | 39.7 | 43 | 42.3 | 43.8 | 43.6 | 42.7 | 44.55 | 54 | 44.881 | 16
| 1 | 4 | 45.3 | 40.2595 | 45.8427 | 48 | 39.6 | 44.7 | 43.2 | 45 | 43.9 | 47.1 | 43.426 | 39.7 | 43 | 42.6 | 44.9 | 43.4 | 43.6 | 43.82 | 52 | 44.223 | 44.0176
| 1 | 5 | 44.4 | 40.327 | 46.0738 | 49 | 40.9 | 43.6 | 41 | 44.4 | 43.5 | 44 | 43.7216 | 39.5 | 43 | 46.8 | 44.1 | 48.4 | 46.2 | 44.99 | 54 | 43.085 | 48
| 1 | 6 | 45 | 39.9197 | 47.2756 | 49 | 39 | 43.9 | 43.7 | 44.2 | 44.5 | 44.5 | 43.6848 | 40.9 | 42 | 43.6 | 44.7 | 44.9 | 43.6 | 43.52 | 54 | 44.661
| Mean |  | 44.7167 | 39.974 | 46.6111 | 48.5 | 40.6 | 43.7 | 42.9667 | 44.15 | 44.55 | 44.8 | 42.9898 | 39.85 | 42.8333 | 42.5 | 44.4667 | 44.8 | 43.3333 | 44.0717 | 53.6667 | 44.153
| Median |  | 44.9 | 40.0896 | 46.0911 | 48.5 | 40.85 | 43.75 | 43.45 | 44.05 | 44.7 | 44.25 | 43.5528 | 39.7 | 43 | 42.45 | 44.5 | 44.25 | 43.35 | 44.025 | 54 | 44.207
| Std Dev. |  | 0.53821 | 0.37062 | 1.2217 | 0.54772 | 1.1402 | 0.76681 | 1.4024 | 0.501 | 0.75829 | 1.626 | 0.99749 | 0.74766 | 0.40825 | 2.7174 | 0.4761 | 1.845 | 1.7455 | 0.63716 | 0.8165 | 0.63422
| Rel.Std.Dev. |  | 1.204% | 0.9271% | 2.621% | 1.129% | 2.808% | 1.755% | 3.264% | 1.135% | 1.702% | 3.63% | 2.32% | 1.876% | 0.9531% | 6.394% | 1.071% | 4.118% | 4.028% | 1.446% | 1.521% | 1.436%
| PDM3 |  | 1.588% | -9.186% | 5.892% | 10.18% | -7.764% | -0.7216% | -2.388% | 0.3007% | 1.209% | 1.777% | -2.335% | -9.468% | -2.69% | -3.448% | 1.02% | 1.777% | -1.555% | 0.1228% | 21.92% | 0.3076%
| Z-Score (Absolute) |  | 0.43 | 2.94 | 1.78 | 3.12 | 2.49 | 0.29 | 0.81 | 0.03 | 0.31 | 0.49 | 0.8 | 3.03 | 0.91 | 1.14 | 0.25 | 0.49 | 0.55 | 0.03 | 6.79 | 0.03
| Analytical results for P in OREAS 263b (Certified Value 0.029 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.0278 | 0.029185 | 0.024 | 0.0304 | 0.028 | 0.029 | 0.03 | 0.029 | 0.029 | 0.030567 | 0.028 | 0.03 | 0.03 | 0.0311 | 0.028 | 0.0343 | 0.03 | 0.025 | 0.029906 |  | 1
| 1 | 2 | 0.0286 | 0.028982 | 0.028 | 0.0303 | 0.028 | 0.03 | 0.029 | 0.029 | 0.029 | 0.029824 | 0.0287 | 0.03 | 0.031 | 0.0326 | 0.028 | 0.0332 | 0.03 | 0.025 | 0.029971 |  | #N/A
| 1 | 3 | 0.0279 | 0.028954 | 0.027 | 0.0307 | 0.028 | 0.029 | 0.029 | 0.029 | 0.029 | 0.03012 | 0.0284 | 0.029 | 0.03 | 0.0304 | 0.028 | 0.0296 | 0.03 | 0.025 | 0.029589 |  | 16
| 1 | 4 | 0.0281 | 0.029283 | 0.028 | 0.0298 | 0.029 | 0.029 | 0.03 | 0.03 | 0.03 | 0.029483 | 0.0287 | 0.029 | 0.031 | 0.029 | 0.028 | 0.0341 | 0.03 | 0.025 | 0.029388 |  | 0.029386
| 1 | 5 | 0.0282 | 0.029817 | 0.03 | 0.0302 | 0.028 | 0.03 | 0.03 | 0.03 | 0.029 | 0.029214 | 0.0281 | 0.03 | 0.032 | 0.0319 | 0.028 | 0.0331 | 0.03 | 0.025 | 0.030033 |  | 49
| 1 | 6 | 0.0287 | 0.028869 | 0.025 | 0.0298 | 0.029 | 0.029 | 0.03 | 0.029 | 0.03 | 0.030074 | 0.0284 | 0.029 | 0.031 | 0.033 | 0.028 | 0.033 | 0.03 | 0.025 | 0.029851
| Mean |  | 0.028217 | 0.029181 | 0.027 | 0.0302 | 0.028333 | 0.029333 | 0.029667 | 0.029333 | 0.029333 | 0.02988 | 0.028383 | 0.0295 | 0.030833 | 0.031333 | 0.028 | 0.032883 | 0.03 | 0.025 | 0.02979
| Median |  | 0.02815 | 0.029083 | 0.0275 | 0.03025 | 0.028 | 0.029 | 0.03 | 0.029 | 0.029 | 0.029949 | 0.0284 | 0.0295 | 0.031 | 0.0315 | 0.028 | 0.03315 | 0.03 | 0.025 | 0.029878
| Std Dev. |  | 0.0003656 | 0.00034742 | 0.0021909 | 0.00035214 | 0.0005164 | 0.0005164 | 0.0005164 | 0.0005164 | 0.0005164 | 0.0004841 | 0.00029269 | 0.00054772 | 0.00075277 | 0.0014882 | 3.8006e-18 | 0.0016987 | 0 | 3.8006e-18 | 0.00024952
| Rel.Std.Dev. |  | 1.296% | 1.191% | 8.114% | 1.166% | 1.823% | 1.76% | 1.741% | 1.76% | 1.76% | 1.62% | 1.031% | 1.857% | 2.441% | 4.749% | 1.357e-14% | 5.166% | 0% | 1.52e-14% | 0.8376%
| PDM3 |  | -3.98% | -0.6974% | -8.121% | 2.769% | -3.583% | -0.1804% | 0.9539% | -0.1804% | -0.1804% | 1.681% | -3.413% | 0.3868% | 4.924% | 6.626% | -4.718% | 11.9% | 2.088% | -14.93% | 1.372%
| Z-Score (Absolute) |  | 0.87 | 0.12 | 1.82 | 0.67 | 0.78 | 0 | 0.26 | 0 | 0 | 0.43 | 0.74 | 0.13 | 1.17 | 1.56 | 1.04 | 2.76 | 0.52 | 3.37 | 0.35
| Analytical results for Pb in OREAS 263b (Certified Value 19.4 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 19.5 | 20.8789 | 17 | 22.1 | 19.1 | 19.1 | 19 | 17.9 | 18.8 | 18.3552 | 21 | 19 | 17.54 | 20.2 | 19.32 | 17.6 | 20.85 | 20 | 19.371 |  | 1
| 1 | 2 | 19.8 | 19.3691 | 17 | 22.4 | 19.1 | 19.3 | 19 | 18.2 | 20.7 | 18.5943 | 23 | 19 | 16.84 | 19.4 | 18.89 | 18.3 | 20.77 | 19 | 19.039 |  | 16
| 1 | 3 | 20.1 | 20.6697 | 16 | 21 | 19.1 | 19.1 | 19.2 | 18.2 | 21.1 | 18.4767 | 22 | 19 | 18.77 | 19.4 | 19.32 | 17.6 | 20.4 | 19 | 17.512 |  | 16
| 1 | 4 | 20 | 20.4595 | 18 | 22.1 | 19.2 | 19.1 | 19.2 | 17.4 | 23.6 | 19.1107 | 19 | 18 | 18.1 | 19.5 | 18.9 | 17.7 | 20.31 | 20 | 19.29 |  | 19.3605
| 1 | 5 | 20.2 | 19.5848 | 18 | 20.9 | 19 | 17.8 | 19.1 | 17.6 | 19.8 | 18.9728 | 21 | 18 | 19.59 | 19.8 | 19.67 | 18.7 | 20.93 | 20 | 18.533 |  | 50
| 1 | 6 | 20.2 | 19.6285 | 18 | 21.1 | 19.2 | 18.9 | 19.6 | 18 | 21.3 | 18.8345 | 21 | 18 | 19.27 | 22 | 19.35 | 17.7 | 20.5 | 19 | 18.663
| Mean |  | 19.9667 | 20.0984 | 17.3333 | 21.6 | 19.1167 | 18.8833 | 19.1833 | 17.8833 | 20.8833 | 18.724 | 21.1667 | 18.5 | 18.3517 | 20.05 | 19.2417 | 17.9333 | 20.6267 | 19.5 | 18.7347
| Median |  | 20.05 | 20.044 | 17.5 | 21.6 | 19.1 | 19.1 | 19.15 | 17.95 | 20.9 | 18.7144 | 21 | 18.5 | 18.435 | 19.65 | 19.32 | 17.7 | 20.635 | 19.5 | 18.851
| Std Dev. |  | 0.27325 | 0.64534 | 0.8165 | 0.66933 | 0.075277 | 0.54559 | 0.22286 | 0.32506 | 1.6216 | 0.29588 | 1.3292 | 0.54772 | 1.0544 | 1.0035 | 0.29929 | 0.45898 | 0.25696 | 0.54772 | 0.68508
| Rel.Std.Dev. |  | 1.369% | 3.211% | 4.711% | 3.099% | 0.3938% | 2.889% | 1.162% | 1.818% | 7.765% | 1.58% | 6.279% | 2.961% | 5.746% | 5.005% | 1.555% | 2.559% | 1.246% | 2.809% | 3.657%
| PDM3 |  | 3.131% | 3.812% | -10.47% | 11.57% | -1.259% | -2.465% | -0.915% | -7.63% | 7.866% | -3.287% | 9.329% | -4.444% | -5.211% | 3.562% | -0.6137% | -7.371% | 6.54% | 0.7207% | -3.232%
| Z-Score (Absolute) |  | 0.64 | 0.74 | NA | 1.96 | 0.05 | 0.24 | 0 | 1.05 | 1.38 | 0.37 | NA | NA | 0.67 | 0.7 | 0.05 | 1.01 | 1.17 | NA | 0.36
Results from laboratories 1.07, 1.15, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 263b (Certified Value 9.18 ppm) | [21 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 9.2 | 9.2245 | 10.3138 | 9.6 | 7.6 | 9.68 | 8.5 | 8.91 | 9 | [11 empty] | 1
| 1 | 2 | 8.9 | 8.9004 | 10.2368 | 9.6 | 7.1 | 9.83 | 8.5 | 9.3 | 9 | [11 empty] | 33
| 1 | 3 | 9.06 | 8.8107 | 10.3286 | 9.8 | 8.6 | 9.64 | 8.6 | 8.99 | 8.8 | [11 empty] | 16
| 1 | 4 | 9.13 | 8.9792 | 10.2574 | 9.4 | 7.5 | 9.73 | 9 | 8.67 | 8.8 | [11 empty] | 9.184
| 1 | 5 | 9.21 | 9.2605 | 10.3026 | 9.5 | 9.3 | 9.69 | 9.1 | 9.12 | 9.05 | [11 empty] | 51
| 1 | 6 | 9.06 | 9.2579 | 10.2561 | 9.7 | 8.3 | 9.87 | 8.5 | 8.86 | 8.95
| Mean |  | 9.0933 | 9.0722 | 10.2825 | 9.6 | 8.0667 | 9.74 | 8.7 | 8.975 | 8.9333
| Median |  | 9.095 | 9.1019 | 10.28 | 9.6 | 7.95 | 9.71 | 8.55 | 8.95 | 8.975
| Std Dev. |  | 0.11483 | 0.19984 | 0.037232 | 0.14142 | 0.8165 | 0.090774 | 0.27568 | 0.21788 | 0.10801
| Rel.Std.Dev. |  | 1.263% | 2.203% | 0.3621% | 1.473% | 10.12% | 0.932% | 3.169% | 2.428% | 1.209%
| PDM3 |  | -0.9877% | -1.218% | 11.96% | 4.529% | -12.17% | 6.053% | -5.271% | -2.276% | -2.73%
| Z-Score (Absolute) |  | 0.04 | 0 | 2.19 | 0.96 | 1.82 | 1.21 | 0.67 | 0.18 | 0.25
| Analytical results for Pt in OREAS 263b (Indicative Value < 5 ppb) | [21 empty] | Indicative
Pt | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.21 | [19 empty] | ppb
|  |  | 4A*OES/MS | [19 empty] | 2
2
| 1 | 1 | < 5 | [19 empty] | 1
| 1 | 2 | < 5 | [19 empty] | 1
| 1 | 3 | < 5 | [19 empty] | 16
| 1 | 4 | < 5 | [19 empty] | < 5
| 1 | 5 | 5 | [19 empty] | 7
| 1 | 6 | < 5
| Mean |  | 5
| Median |  | 5
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 263b (Certified Value 132 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 131.98 | 129.594 | 161.1475 | 121 | 134 | 139 | 128 | 129 | 124.5 | 128.194 | 139 | 136 | 90.3 | 132.52 | 130.8 | 130 | 125 |  |  |  | 1
| 1 | 2 | 131.74 | 127.8084 | 161.176 | 126 | 133 | 139 | 131 | 124.5 | 130.5 | 126.9482 | 135.2 | 132 | 100.8 | 132.03 | 128.2 | 134 | 127 |  |  |  | 34
| 1 | 3 | 135.77 | 129.9174 | 161.183 | 120 | 136 | 135 | 130 | 122.5 | 140 | 127.3132 | 137.2 | 133 | 112.6 | 131.64 | 126 | 133 | 125 |  |  |  | 16
| 1 | 4 | 134.91 | 129.0605 | 161.141 | 123 | 137.5 | 137.5 | 132.5 | 123.5 | 131 | 130.9224 | 136.7 | 135 | 106.5 | 131.93 | 128.7 | 130 | 125 |  |  |  | 131.5962
| 1 | 5 | 133.7 | 131.1842 | 161.1995 | 121 | 136.5 | 126 | 131 | 124 | 131.5 | 130.9055 | 138.9 | 133 | 118.1 | 133.64 | 130 | 138 | 127 |  |  |  | 52
| 1 | 6 | 131.66 | 127.0047 | 161.257 | 124 | 138 | 138 | 135.5 | 126 | 131 | 129.6993 | 139.9 | 132 | 116.6 | 132.51 | 133 | 130 | 126
| Mean |  | 133.2933 | 129.0949 | 161.184 | 122.5 | 135.8333 | 135.75 | 131.3333 | 124.9167 | 131.4167 | 128.9971 | 137.8167 | 133.5 | 107.4833 | 132.3783 | 129.45 | 132.5 | 125.8333
| Median |  | 132.84 | 129.3272 | 161.1795 | 122 | 136.25 | 137.75 | 131 | 124.25 | 131 | 128.9466 | 138.05 | 133 | 109.55 | 132.27 | 129.35 | 131.5 | 125.5
| Std Dev. |  | 1.773 | 1.5044 | 0.041989 | 2.2583 | 1.9664 | 4.9975 | 2.5232 | 2.3112 | 4.954 | 1.7618 | 1.7543 | 1.6432 | 10.6105 | 0.70664 | 2.3981 | 3.2094 | 0.98319
| Rel.Std.Dev. |  | 1.33% | 1.165% | 0.02605% | 1.844% | 1.448% | 3.681% | 1.921% | 1.85% | 3.77% | 1.366% | 1.273% | 1.231% | 9.872% | 0.5338% | 1.853% | 2.422% | 0.7813%
| PDM3 |  | 1.29% | -1.901% | 22.48% | -6.912% | 3.22% | 3.156% | -0.1998% | -5.076% | -0.1364% | -1.975% | 4.727% | 1.447% | -18.32% | 0.5943% | -1.631% | 0.6868% | -4.379%
| Z-Score (Absolute) |  | 0.52 | 0.65 | 8.3 | 2.48 | 1.23 | 1.21 | 0.02 | 1.81 | 0 | 0.67 | 1.78 | 0.58 | 6.67 | 0.27 | 0.55 | 0.3 | 1.56
| Analytical results for Re in OREAS 263b (Indicative Value 0.001 ppm) | [21 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 3
3
| 1 | 1 | 0.013007 | < 0.001 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.05 | 0.015 | < 0.05 | < 0.002 | < 0.002 | 0.0012 | < 0.1 | [6 empty] | 1
| 1 | 2 | 0.01267 | < 0.001 | 0.002 | < 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | 0.014 | < 0.05 | < 0.002 | < 0.002 | 0.0006 | < 0.1 | [6 empty] | 2
| 1 | 3 | 0.013272 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.002 | < 0.05 | 0.016 | < 0.05 | 0.002 | < 0.002 | 0.0018 | < 0.1 | [6 empty] | 16
| 1 | 4 | 0.013413 | < 0.001 | 0.002 | < 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | 0.014 | < 0.05 | < 0.002 | < 0.002 | 0.0015 | < 0.1 | [6 empty] | 0.0014313
| 1 | 5 | 0.012581 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | 0.012 | < 0.05 | 0.003 | < 0.002 | 0.001 | < 0.1 | [6 empty] | 8
| 1 | 6 | 0.012613 | < 0.001 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.002 | < 0.05 | 0.014 | < 0.05 | < 0.002 | < 0.002 | 0.0011 | < 0.1
| Mean |  | 0.012926 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 |  | 0.014167 |  | 0.0025 |  | 0.0012
| Median |  | 0.012838 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 |  | 0.014 |  | 0.0025 |  | 0.00115
| Std Dev. |  | 0.00035919 |  | 0 | 0 | 0 | 0 | 0 |  | 0.0013292 |  | 0.00070711 |  | 0.00041473
| Rel.Std.Dev. |  | 2.779% |  | 0% | 0% | 0% | 0% | 0% |  | 9.382% |  | 28.28% |  | 34.56%
| PDM3 |  | 803.1% | 39.74% | 39.74% | 39.74% | 39.74% | 39.74% | 39.74% |  | 889.8% |  | 74.67% |  | -16.16%
| Z-Score (Absolute) |  | 15.3 | 1.12 | 0.11 | 0.34 | 0.11 | 0.56 | 0.11 | 31.58 | 16.97 | 31.58 | 0.11 | 0.79 | 0.52 | 65.3
| Analytical results for S in OREAS 263b (Certified Value 0.029 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.026 | 0.033153 | 0.0259 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027359 | 0.0268 | 0.031 | < 0.04 | < 0.04 | 0.0201 | 0.02 | 0.035 | 0.02267 |  | 1
| 1 | 2 | 0.0262 | 0.03745 | 0.026034 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.027344 | 0.0265 | 0.03 | < 0.04 | < 0.04 | 0.0223 | 0.02 | 0.035 | 0.021112 |  | 20
| 1 | 3 | 0.0265 | 0.037815 | 0.026996 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027209 | 0.0259 | 0.031 | < 0.04 | < 0.04 | 0.0252 | 0.01 | 0.035 | 0.018802 |  | 16
| 1 | 4 | 0.0267 | 0.03486 | 0.0259 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.027075 | 0.0262 | 0.031 | < 0.04 | < 0.04 | 0.0251 | 0.02 | 0.035 | 0.01086 |  | 0.028769
| 1 | 5 | 0.0263 | 0.043055 | 0.0278 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026584 | 0.0273 | 0.03 | < 0.04 | < 0.04 | 0.0179 | 0.02 | 0.04 | 0.024143 |  | 53
| 1 | 6 | 0.0261 | 0.041497 | 0.026868 | 0.02 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026542 | 0.0273 | 0.03 | < 0.04 | < 0.04 | 0.0224 | 0.02 | 0.04 | 0.022607
| Mean |  | 0.0263 | 0.037972 | 0.026583 | 0.028333 | 0.031667 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033333 | 0.027019 | 0.026667 | 0.0305 |  |  | 0.022167 | 0.018333 | 0.036667 | 0.020032
| Median |  | 0.02625 | 0.037633 | 0.026451 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027142 | 0.02665 | 0.0305 |  |  | 0.02235 | 0.02 | 0.035 | 0.021859
| Std Dev. |  | 0.00026077 | 0.0037808 | 0.00077033 | 0.0040825 | 0.0040825 | 0 | 0 | 0 | 0 | 0.005164 | 0.00036811 | 0.00057504 | 0.00054772 |  |  | 0.0028409 | 0.0040825 | 0.002582 | 0.0048434
| Rel.Std.Dev. |  | 0.9915% | 9.957% | 2.898% | 14.41% | 12.89% | 0% | 0% | 0% | 0% | 15.49% | 1.362% | 2.156% | 1.796% |  |  | 12.82% | 22.27% | 7.042% | 24.18%
| PDM3 |  | -8.582% | 31.99% | -7.598% | -1.515% | 10.07% | 4.279% | 4.279% | 4.279% | 4.279% | 15.87% | -6.083% | -7.308% | 6.017% |  |  | -22.95% | -36.27% | 27.45% | -30.37%
| Z-Score (Absolute) |  | 0.63 | 3 | 0.54 | 0 | 1.04 | 0.52 | 0.52 | 0.52 | 0.52 | 1.56 | 0.41 | 0.52 | 0.67 | 2.59 | 2.59 | 1.92 | 3.11 | 2.59 | 2.58
| Analytical results for Sb in OREAS 263b (Certified Value 11.3 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 11.24 | 11.1933 | 10.8 | 8.38 | 11.4 | 12.65 | 11.3 | 11.6 | 11.35 | 11.2337 | 10.95 | 10.6 | 10.24 | 11.83 | 9.71 | 12.6 | 13.88 | 11.9 | 14.331 |  | 1
| 1 | 2 | 11.09 | 11.1088 | 10.9 | 8.98 | 11.35 | 12.7 | 11.1 | 11.5 | 12.1 | 10.8241 | 10.71 | 10 | 9.93 | 11.66 | 9.67 | 12.9 | 12.19 | 12 | 13.077 |  | 36
| 1 | 3 | 11.44 | 11.221 | 10.5 | 9.69 | 11.75 | 11.9 | 11 | 12.1 | 13.05 | 10.3726 | 10.81 | 10.1 | 10.4 | 11.67 | 9.64 | 12.2 | 14.81 | 11.9 | 10.749 |  | 16
| 1 | 4 | 11.09 | 11.1423 | 10.9 | 9.57 | 11.65 | 12.35 | 11.35 | 11.2 | 12.75 | 11.1821 | 10.65 | 10.2 | 10.44 | 11.66 | 9.5 | 12.2 | 13.42 | 11.8 | 11.602 |  | 11.3175
| 1 | 5 | 11.23 | 11.5533 | 11.2 | 9.37 | 11.65 | 11.5 | 11.3 | 11.3 | 11.95 | 10.5375 | 10.8 | 10.6 | 11 | 11.59 | 9.79 | 13 | 12.77 | 12 | 11.244 |  | 54
| 1 | 6 | 10.88 | 11.0516 | 11 | 8.62 | 11.65 | 13.05 | 11.5 | 11.4 | 12.4 | 11.5033 | 11.01 | 10.3 | 10.9 | 12.04 | 9.86 | 12.6 | 13.38 | 12 | 12.173
| Mean |  | 11.1617 | 11.2117 | 10.8833 | 9.1017 | 11.575 | 12.3583 | 11.2583 | 11.5167 | 12.2667 | 10.9422 | 10.8217 | 10.3 | 10.485 | 11.7417 | 9.695 | 12.5833 | 13.4083 | 11.9333 | 12.196
| Median |  | 11.16 | 11.1678 | 10.9 | 9.175 | 11.65 | 12.5 | 11.3 | 11.45 | 12.25 | 11.0031 | 10.805 | 10.25 | 10.42 | 11.665 | 9.69 | 12.6 | 13.4 | 11.95 | 11.8875
| Std Dev. |  | 0.18862 | 0.17783 | 0.23166 | 0.53011 | 0.16047 | 0.57045 | 0.18005 | 0.31885 | 0.60553 | 0.43806 | 0.13747 | 0.25298 | 0.40367 | 0.1663 | 0.12502 | 0.33714 | 0.90331 | 0.08165 | 1.3182
| Rel.Std.Dev. |  | 1.69% | 1.586% | 2.129% | 5.824% | 1.386% | 4.616% | 1.599% | 2.769% | 4.936% | 4.003% | 1.27% | 2.456% | 3.85% | 1.416% | 1.29% | 2.679% | 6.737% | 0.6842% | 10.81%
| PDM3 |  | -1.377% | -0.9344% | -3.836% | -19.58% | 2.276% | 9.197% | -0.5224% | 1.76% | 8.387% | -3.315% | -4.381% | -8.99% | -7.355% | 3.748% | -14.34% | 11.19% | 18.47% | 5.442% | 7.763%
| Z-Score (Absolute) |  | 0.1 | 0.05 | 0.37 | 2.15 | 0.32 | 1.1 | 0 | 0.26 | 1.01 | 0.32 | 0.44 | 0.96 | 0.77 | 0.48 | 1.56 | 1.32 | 2.15 | 0.67 | 0.94
| Analytical results for Sc in OREAS 263b (Certified Value 10.9 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 11.4 | 11.1333 | 13.551 | 12 | 11 | 11.4 | 10.6 | 11 | 11.4 | 10.2 | 11.8202 | 10.33 | 11 | 11.6 | 11.3 | 10.4 | 12 | 11 | 10.526 |  | 1
| 1 | 2 | 11.2 | 10.9714 | 13.546 | 12 | 11 | 11.6 | 10.7 | 11.2 | 11 | 10.5 | 11.494 | 10.38 | 11 | 9.3 | 11.2 | 10.1 | 11 | 11 | 10.65 |  | 37
| 1 | 3 | 11.4 | 10.7864 | 13.626 | 12 | 11 | 11.8 | 10.1 | 11.2 | 11.2 | 11.4 | 11.3194 | 10.19 | 11 | 10.7 | 11 | 9.9 | 11 | 12 | 10.489 |  | 16
| 1 | 4 | 11.5 | 10.9539 | 13.585 | 13 | 11 | 11.6 | 10.3 | 11.4 | 10.8 | 10.7 | 11.4251 | 10.19 | 11 | 9.6 | 11.4 | 10.2 | 12 | 11 | 10.52 |  | 10.9153
| 1 | 5 | 11.6 | 11.1185 | 13.512 | 13 | 11 | 11.7 | 9.4 | 11.2 | 10.8 | 10.7 | 11.9584 | 10.19 | 11 | 11.2 | 11.3 | 10.4 | 13 | 11 | 10.671 |  | 55
| 1 | 6 | 11.2 | 11.0688 | 13.497 | 12 | 11 | 11.8 | 10.5 | 11.5 | 10.9 | 10.8 | 11.9314 | 10.37 | 11 | 10.5 | 11.5 | 10.4 | 12 | 11 | 10.697
| Mean |  | 11.3833 | 11.0054 | 13.5528 | 12.3333 | 11 | 11.65 | 10.2667 | 11.25 | 11.0167 | 10.7167 | 11.6581 | 10.275 | 11 | 10.4833 | 11.2833 | 10.2333 | 11.8333 | 11.1667 | 10.5922
| Median |  | 11.4 | 11.0201 | 13.5485 | 12 | 11 | 11.65 | 10.4 | 11.2 | 10.95 | 10.7 | 11.6571 | 10.26 | 11 | 10.6 | 11.3 | 10.3 | 12 | 11 | 10.588
| Std Dev. |  | 0.16021 | 0.13032 | 0.047377 | 0.5164 | 0 | 0.15166 | 0.4761 | 0.17607 | 0.24014 | 0.39707 | 0.27824 | 0.094604 | 0 | 0.89312 | 0.17224 | 0.20656 | 0.75277 | 0.40825 | 0.090309
| Rel.Std.Dev. |  | 1.407% | 1.184% | 0.3496% | 4.187% | 0% | 1.302% | 4.637% | 1.565% | 2.18% | 3.705% | 2.387% | 0.9207% | 0% | 8.519% | 1.527% | 2.018% | 6.361% | 3.656% | 0.8526%
| PDM3 |  | 4.288% | 0.8255% | 24.16% | 12.99% | 0.7761% | 6.731% | -5.942% | 3.067% | 0.9288% | -1.82% | 6.805% | -5.866% | 0.7761% | -3.957% | 3.372% | -6.248% | 8.411% | 2.303% | -2.96%
| Z-Score (Absolute) |  | 0.62 | 0 | 4.16 | NA | 0.01 | 1.05 | 1.21 | 0.4 | 0.02 | 0.47 | 1.07 | 1.19 | NA | 0.85 | 0.45 | 1.26 | NA | NA | 0.67
Results from laboratories 1.07, 1.16, 1.21 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 263b (Indicative Value 0.86 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.5 | < 1.5 | 0.5 | 1 | 1 | < 1 | < 1 | 2 | < 1 | < 2 | 0.48 | < 1 | 0.6 |  | 0.5 | 1.88 | < 5 | 6.503 |  |  | 1
| 1 | 2 |  | < 1.5 | 0.7 | 1 | 1 | < 1 | 1 | 1 | 1 | < 2 | 0.47 | < 1 | 1 | 0.5 | 0.5 | 1.73 | < 5 | 5.491 |  |  | 38
| 1 | 3 |  | < 1.5 | 0.5 | 1 | 1 | 1 | < 1 | 1 | 1 | < 2 | 0.48 | < 1 | 0.3 |  | 0.6 | 1.77 | < 5 | 4.644 |  |  | 16
| 1 | 4 |  | < 1.5 | 0.5 | 1 | 1 | 1 | 1 | 1 | 1 | < 2 | 0.47 | < 1 | 0.6 |  | 0.6 | 1.24 | < 5 | 6.689 |  |  | 0.86026
| 1 | 5 | 0.6 | < 1.5 | 0.6 | 1 | 1 | 1 | 1 | 1 | 1 | < 2 | 0.47 | < 1 | 0.8 |  | 0.4 | 1.25 | < 5 |  |  |  | 7
| 1 | 6 | 0.5 | < 1.5 | 0.4 | 1 | 1 | 1 | < 1 | 2 | 1 | < 2 | 0.45 | < 1 | 0.8 |  | 0.7 | 1.61 | < 5 | 3.669
| Mean |  | 0.53333 |  | 0.53333 | 1 | 1 | 1 | 1 | 1.3333 | 1 |  | 0.47 |  | 0.68333 | 0.5 | 0.55 | 1.58 |  | 5.3992
| Median |  | 0.5 |  | 0.5 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.47 |  | 0.7 | 0.5 | 0.55 | 1.67 |  | 5.491
| Std Dev. |  | 0.057735 |  | 0.10328 | 0 | 0 | 0 | 0 | 0.5164 | 0 |  | 0.010954 |  | 0.24014 |  | 0.10488 | 0.2735 |  | 1.2703
| Rel.Std.Dev. |  | 10.83% |  | 19.36% | 0% | 0% | 0% | 0% | 38.73% | 0% |  | 2.331% |  | 35.14% |  | 19.07% | 17.31% |  | 23.53%
| PDM3 |  | -38% |  | -38% | 16.24% | 16.24% | 16.24% | 16.24% | 54.99% | 16.24% |  | -45.37% |  | -20.57% | -41.88% | -36.07% | 83.67% |  | 527.6%
| Z-Score (Absolute) |  | 0.7 | 0.11 | 0.7 | 0.56 | 0.56 | 0.11 | 0.11 | 1.46 | 0.34 | 0.56 | 0.87 | 0.79 | 0.29 | 0.79 | 0.65 | 2.13 | 4.61 | 12.43
| Analytical results for Sm in OREAS 263b (Certified Value 5.86 ppm) | [21 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 5.91 | 6.0985 | 6.5354 | 5.9 | 5.4 | 6.04 | 5.5 | 5.66 | 6.05 | [11 empty] | 1
| 1 | 2 | 5.62 | 5.6084 | 6.4226 | 6 | 4.9 | 6.18 | 5.4 | 5.8 | 6.05 | [11 empty] | 5
| 1 | 3 | 6.03 | 5.6469 | 6.5758 | 5.8 | 5.5 | 6.09 | 5.3 | 5.33 | 6.1 | [11 empty] | 16
| 1 | 4 | 5.82 | 5.6828 | 6.5277 | 6.1 | 5.2 | 6.03 | 5.6 | 5.38 | 6.05 | [11 empty] | 5.8589
| 1 | 5 | 5.78 | 5.8899 | 6.5208 | 6.1 | 6 | 6.04 | 5.7 | 5.64 | 5.85 | [11 empty] | 56
| 1 | 6 | 5.84 | 5.9655 | 6.5553 | 5.8 | 5.4 | 6.13 | 5.5 | 5.51 | 6.1
| Mean |  | 5.8333 | 5.8153 | 6.5229 | 5.95 | 5.4 | 6.085 | 5.5 | 5.5533 | 6.0333
| Median |  | 5.83 | 5.7864 | 6.5316 | 5.95 | 5.4 | 6.065 | 5.5 | 5.575 | 6.05
| Std Dev. |  | 0.13648 | 0.1985 | 0.053114 | 0.13784 | 0.36332 | 0.060249 | 0.14142 | 0.17974 | 0.093095
| Rel.Std.Dev. |  | 2.34% | 3.413% | 0.8143% | 2.317% | 6.728% | 0.9901% | 2.571% | 3.237% | 1.543%
| PDM3 |  | -0.4361% | -0.7433% | 11.33% | 1.555% | -7.832% | 3.859% | -6.125% | -5.215% | 2.978%
| Z-Score (Absolute) |  | 0 | 0.05 | 1.85 | 0.31 | 1.16 | 0.67 | 0.89 | 0.75 | 0.54
| Analytical results for Sn in OREAS 263b (Certified Value 11.6 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 11.9 | 11.4539 | 2 | 11.5 | 12.4 | 12 | 10.8 | 12.6 | 12 | 11.1318 | 8.98 | 11.2 | 10.9 | 12.4 | 10.3 | 14 | < 5 | 12.7 | 11.376 |  | 1
| 1 | 2 | 11.6 | 11.5492 | 3 | 11.2 | 12.4 | 12.1 | 11.1 | 12.2 | 12.5 | 10.9806 | 9.48 | 10.7 | 10.7 | 12.2 | 10.5 | 13.9 | < 5 | 12.7 | 12.08 |  | 11
| 1 | 3 | 11.9 | 11.2452 | 3 | 11.4 | 12.5 | 11.3 | 11.2 | 12.2 | 13.3 | 11.1655 | 9.18 | 10.8 | 11.5 | 12.3 | 10.4 | 13.4 | < 5 | 12.6 | 12.348 |  | 16
| 1 | 4 | 11.7 | 11.3777 | 5 | 11.4 | 12.3 | 11.4 | 11.3 | 12 | 15.6 | 11.3575 | 8.98 | 10.5 | 11.5 | 12.2 | 10.3 | 13.7 | < 5 | 12.7 | 12.035 |  | 11.6399
| 1 | 5 | 11.7 | 11.3694 | 4 | 10.9 | 12.4 | 10.8 | 11 | 11.8 | 12.2 | 10.9732 | 8.88 | 10.5 | 12.4 | 12.1 | 10.7 | 14 | < 5 | 12.8 | 11.132 |  | 57
| 1 | 6 | 11.6 | 11.4103 | 2 | 11 | 12.5 | 11.7 | 11.2 | 12.6 | 12.5 | 11.4347 | 9.38 | 10.9 | 11.8 | 12.4 | 10.6 | 13.7 | < 5 | 12.7 | 12.672
| Mean |  | 11.7333 | 11.401 | 3.1667 | 11.2333 | 12.4167 | 11.55 | 11.1 | 12.2333 | 13.0167 | 11.1739 | 9.1467 | 10.7667 | 11.4667 | 12.2667 | 10.4667 | 13.7833 |  | 12.7 | 11.9405
| Median |  | 11.7 | 11.394 | 3 | 11.3 | 12.4 | 11.55 | 11.15 | 12.2 | 12.5 | 11.1486 | 9.08 | 10.75 | 11.5 | 12.25 | 10.45 | 13.8 |  | 12.7 | 12.0575
| Std Dev. |  | 0.13663 | 0.10066 | 1.169 | 0.24221 | 0.075277 | 0.48477 | 0.17889 | 0.32042 | 1.3408 | 0.19039 | 0.24221 | 0.26583 | 0.61536 | 0.12111 | 0.1633 | 0.23166 |  | 0.063246 | 0.58342
| Rel.Std.Dev. |  | 1.164% | 0.8829% | 36.92% | 2.156% | 0.6063% | 4.197% | 1.612% | 2.619% | 10.3% | 1.704% | 2.648% | 2.469% | 5.367% | 0.9873% | 1.56% | 1.681% |  | 0.498% | 4.886%
| PDM3 |  | 0.8025% | -2.053% | -72.79% | -3.493% | 6.673% | -0.7725% | -4.639% | 5.098% | 11.83% | -4.004% | -21.42% | -7.502% | -1.488% | 5.384% | -10.08% | 18.41% |  | 9.107% | 2.582%
| Z-Score (Absolute) |  | 0.21 | 0.1 | NA | 0.25 | 0.84 | 0.04 | 0.38 | 0.67 | 1.39 | 0.31 | 2.17 | 0.68 | 0.04 | 0.7 | 0.96 | 2.09 | 8.28 | 1.1 | 0.4
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 263b (Certified Value 52 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
1
| 1 | 1 | 51.98 | 50.3864 | 52.848 | 56.9 | 46.9 | 52.9 | 52.9 | 53.9 | 54.4 | 51.3 | 52.3889 | 44 | 51.6 | 55 | 53.25 | 50 | 44 | 54.74 | 53 | 48.388 | 1
| 1 | 2 | 52.87 | 49.9144 | 52.4134 | 55.5 | 50 | 52.8 | 53.6 | 55.4 | 52.7 | 53.1 | 50.6335 | 43 | 50.4 | 56 | 53.64 | 50 | 43.1 | 53.93 | 53 | 49.308 | 12
| 1 | 3 | 55.12 | 49.498 | 54.5546 | 57.2 | 49.5 | 53.7 | 51 | 55 | 51.8 | 53.6 | 50.6711 | 44 | 50.4 | 52 | 53.83 | 49 | 39.8 | 54.38 | 53 | 49.34 | 16
| 1 | 4 | 53.45 | 49.9585 | 53.6112 | 57.4 | 47.3 | 53.8 | 52.1 | 54.6 | 51.6 | 52.5 | 53.1078 | 43 | 49 | 61 | 53.27 | 50 | 43.2 | 54.81 | 52.5 | 48.431 | 52.4419
| 1 | 5 | 53.35 | 50.7989 | 54.332 | 57.6 | 47.4 | 53.9 | 48.5 | 55.6 | 52.6 | 52.7 | 52.3606 | 43 | 49.4 | 54 | 53.42 | 51 | 46 | 53.26 | 53 | 48.802 | 58
| 1 | 6 | 52.74 | 49.708 | 54.8302 | 57.1 | 44.4 | 53.9 | 53.3 | 57.1 | 53 | 51.5 | 54.0602 | 45 | 51.2 | 51 | 54.64 | 51 | 42.5 | 53.81 | 53 | 48.967
| Mean |  | 53.2517 | 50.0441 | 53.7649 | 56.95 | 47.5833 | 53.5 | 51.9 | 55.2667 | 52.6833 | 52.45 | 52.2037 | 43.6667 | 50.3333 | 54.8333 | 53.675 | 50.1667 | 43.1 | 54.155 | 52.9167 | 48.8727
| Median |  | 53.11 | 49.9365 | 53.9716 | 57.15 | 47.35 | 53.75 | 52.5 | 55.2 | 52.65 | 52.6 | 52.3747 | 43.5 | 50.4 | 54.5 | 53.53 | 50 | 43.15 | 54.155 | 53 | 48.8845
| Std Dev. |  | 1.0549 | 0.47377 | 0.97683 | 0.75033 | 2.0114 | 0.5099 | 1.911 | 1.0838 | 1.0008 | 0.89833 | 1.3518 | 0.8165 | 1.0013 | 3.5449 | 0.52271 | 0.75277 | 2.0239 | 0.59876 | 0.20412 | 0.41274
| Rel.Std.Dev. |  | 1.981% | 0.9467% | 1.817% | 1.318% | 4.227% | 0.9531% | 3.682% | 1.961% | 1.9% | 1.713% | 2.59% | 1.87% | 1.989% | 6.465% | 0.9739% | 1.501% | 4.696% | 1.106% | 0.3857% | 0.8445%
| PDM3 |  | 1.544% | -4.572% | 2.523% | 8.596% | -9.265% | 2.018% | -1.033% | 5.387% | 0.4604% | 0.01551% | -0.4542% | -16.73% | -4.021% | 4.56% | 2.351% | -4.339% | -17.81% | 3.267% | 0.9054% | -6.806%
| Z-Score (Absolute) |  | 0.24 | 0.89 | 0.42 | 1.55 | 1.76 | 0.33 | 0.24 | 0.95 | 0.04 | 0.04 | 0.13 | 3.14 | 0.79 | 0.8 | 0.39 | 0.85 | 3.34 | 0.56 | 0.12 | 1.3
| Analytical results for Ta in OREAS 263b (Certified Value 0.93 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.86 | 1.0209 | < 0.1 | 0.53 | 0.91 | 0.92 | 0.96 | 0.91 | 0.93 | 0.15752 | 0.45 | 0.6 | 0.9 | 0.86 | 0.7 | 1 | 1 |  |  |  | 1
| 1 | 2 | 0.83 | 0.95916 | < 0.1 | 0.54 | 0.91 | 0.96 | 0.93 | 0.91 | 0.93 | 0.22258 | 0.45 | 0.52 | 0.8 | 0.88 | 0.7 | 1.01 | 1 |  |  |  | 25
| 1 | 3 | 0.92 | 0.93531 | < 0.1 | 0.53 | 0.95 | 0.91 | 0.94 | 0.86 | 0.93 | 0.18903 | 0.44 | 0.54 | 0.8 | 0.93 | 0.7 | 0.93 | 1 |  |  |  | 16
| 1 | 4 | 0.87 | 0.9113 | < 0.1 | 0.51 | 0.9 | 0.95 | 0.95 | 0.89 | 0.95 | 0.18655 | 0.44 | 0.49 | 1 | 0.89 | 0.7 | 0.96 | 1 |  |  |  | 0.92955
| 1 | 5 | 0.98 | 1.0108 | < 0.1 | 0.54 | 0.92 | 0.84 | 0.93 | 0.94 | 0.98 | 0.15546 | 0.45 | 0.47 | 0.9 | 0.85 | 0.7 | 1.02 | 1 |  |  |  | 59
| 1 | 6 | 0.91 | 0.95405 | < 0.1 | 0.5 | 0.94 | 0.96 | 0.97 | 0.93 | 0.93 | 0.23218 | 0.44 | 0.6 | 0.9 | 0.92 | 0.7 | 0.99 | 1
| Mean |  | 0.895 | 0.96527 |  | 0.525 | 0.92167 | 0.92333 | 0.94667 | 0.90667 | 0.94167 | 0.19055 | 0.445 | 0.53667 | 0.88333 | 0.88833 | 0.7 | 0.985 | 1
| Median |  | 0.89 | 0.95661 |  | 0.53 | 0.915 | 0.935 | 0.945 | 0.91 | 0.93 | 0.18779 | 0.445 | 0.53 | 0.9 | 0.885 | 0.7 | 0.995 | 1
| Std Dev. |  | 0.053198 | 0.042776 |  | 0.016432 | 0.019408 | 0.045898 | 0.01633 | 0.028752 | 0.020412 | 0.031937 | 0.0054772 | 0.05465 | 0.075277 | 0.031885 | 1.2162e-16 | 0.033912 | 0
| Rel.Std.Dev. |  | 5.944% | 4.431% |  | 3.13% | 2.106% | 4.971% | 1.725% | 3.171% | 2.168% | 16.76% | 1.231% | 10.18% | 8.522% | 3.589% | 1.737e-14% | 3.443% | 0%
| PDM3 |  | -3.717% | 3.842% |  | -43.52% | -0.8479% | -0.6686% | 1.842% | -2.462% | 1.304% | -79.5% | -52.13% | -42.27% | -4.972% | -4.434% | -24.69% | 5.965% | 7.579%
| Z-Score (Absolute) |  | 0.07 | 0.79 | 10.41 | 4.6 | 0.25 | 0.28 | 0.56 | 0.07 | 0.5 | 8.69 | 5.58 | 4.45 | NA | 0.15 | NA | 1.03 | NA
| Analytical results for Tb in OREAS 263b (Certified Value 0.56 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
2
| 1 | 1 | 0.5 | 0.59702 | 0.71441 | 0.7 | 0.5 | 0.55 | 0.5 | 0.57 | 0.64 | 0.502 | [10 empty] | 1
| 1 | 2 | 0.54 | 0.54403 | 0.70433 | 0.6 | 0.5 | 0.56 | 0.6 | 0.59 | 0.6 | 1.089 | [10 empty] | 26
| 1 | 3 | 0.54 | 0.55874 | 0.70955 | 0.6 | 0.5 | 0.57 | 0.5 | 0.59 | 0.64 | 1.045 | [10 empty] | 16
| 1 | 4 | 0.51 | 0.56619 | 0.72728 | 0.6 | 0.5 | 0.56 | 0.5 | 0.6 | 0.64 | 0.369 | [10 empty] | 0.56329
| 1 | 5 | 0.5 | 0.57259 | 0.73547 | 0.6 | 0.6 | 0.56 | 0.5 | 0.61 | 0.6 | 0.281 | [10 empty] | 60
| 1 | 6 | 0.52 | 0.58502 | 0.71873 | 0.7 | 0.5 | 0.56 | 0.5 | 0.57 | 0.64 | 1.011
| Mean |  | 0.51833 | 0.5706 | 0.7183 | 0.63333 | 0.51667 | 0.56 | 0.51667 | 0.58833 | 0.62667 | 0.71617
| Median |  | 0.515 | 0.56939 | 0.71657 | 0.6 | 0.5 | 0.56 | 0.5 | 0.59 | 0.64 | 0.7565
| Std Dev. |  | 0.018348 | 0.018846 | 0.011511 | 0.05164 | 0.040825 | 0.0063246 | 0.040825 | 0.016021 | 0.020656 | 0.37144
| Rel.Std.Dev. |  | 3.54% | 3.303% | 1.603% | 8.154% | 7.902% | 1.129% | 7.902% | 2.723% | 3.296% | 51.86%
| PDM3 |  | -7.981% | 1.298% | 27.52% | 12.43% | -8.277% | -0.584% | -8.277% | 4.446% | 11.25% | 27.14%
| Z-Score (Absolute) |  | 0.72 | 0.1 | 1.63 | 0.63 | 0.74 | 0.23 | 0.74 | 0.1 | 0.55 | 1.6
| Analytical results for Te in OREAS 263b (Certified Value 0.26 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.2 | 0.20858 | < 0.1 | 0.48 | 0.27 | 0.25 | 0.23 | 0.33 | 0.25 | 0.25122 | 0.29 | 0.3 | 0.28 | 0.2 | 0.3 | 0.27 | 0.4 |  |  |  | 1
| 1 | 2 | 0.3 | 0.27378 | < 0.1 | 0.49 | 0.28 | 0.26 | 0.21 | 0.3 | 0.25 | 0.26095 | 0.29 | 0.2 | 0.27 | 0.3 | 0.27 | 0.28 | 0.4 |  |  |  | 27
| 1 | 3 | 0.3 | 0.21404 | < 0.1 | 0.52 | 0.24 | 0.22 | 0.18 | 0.31 | 0.29 | 0.27262 | 0.26 | 0.3 | 0.3 | 0.2 | 0.28 | 0.24 | 0.4 |  |  |  | 16
| 1 | 4 | 0.3 | 0.20967 | < 0.1 | 0.47 | 0.27 | 0.26 | 0.27 | 0.32 | 0.23 | 0.24738 | 0.25 | 0.2 | 0.27 | 0.2 | 0.27 | 0.26 | 0.4 |  |  |  | 0.25997
| 1 | 5 | 0.2 | 0.2197 | < 0.1 | 0.52 | 0.25 | 0.26 | 0.16 | 0.34 | 0.27 | 0.26479 | 0.21 | 0.3 | 0.3 | 0.2 | 0.28 | 0.27 | 0.4 |  |  |  | 61
| 1 | 6 | 0.3 | 0.21326 | < 0.1 | 0.53 | 0.24 | 0.25 | 0.23 | 0.29 | 0.25 | 0.23702 | 0.25 | 0.3 | 0.25 | 0.2 | 0.31 | 0.27 | 0.4
| Mean |  | 0.26667 | 0.22317 |  | 0.50167 | 0.25833 | 0.25 | 0.21333 | 0.315 | 0.25667 | 0.25566 | 0.25833 | 0.26667 | 0.27833 | 0.21667 | 0.285 | 0.265 | 0.4
| Median |  | 0.3 | 0.21365 |  | 0.505 | 0.26 | 0.255 | 0.22 | 0.315 | 0.25 | 0.25608 | 0.255 | 0.3 | 0.275 | 0.2 | 0.28 | 0.27 | 0.4
| Std Dev. |  | 0.05164 | 0.025101 |  | 0.024833 | 0.017224 | 0.015492 | 0.039328 | 0.018708 | 0.020656 | 0.01292 | 0.029944 | 0.05164 | 0.019408 | 0.040825 | 0.016432 | 0.013784 | 6.0809e-17
| Rel.Std.Dev. |  | 19.36% | 11.25% |  | 4.95% | 6.667% | 6.197% | 18.43% | 5.939% | 8.048% | 5.054% | 11.59% | 19.36% | 6.973% | 18.84% | 5.766% | 5.202% | 1.52e-14%
| PDM3 |  | 2.574% | -14.16% |  | 92.97% | -0.6311% | -3.837% | -17.94% | 21.17% | -1.272% | -1.659% | -0.6311% | 2.574% | 7.062% | -16.66% | 9.626% | 1.933% | 53.86%
| Z-Score (Absolute) |  | NA | 1.19 | 7.02 | 8.2 | 0 | 0.28 | 1.52 | 1.91 | 0.06 | 0.09 | 0 | NA | 0.67 | NA | 0.9 | 0.22 | NA
Results from laboratories 1.02, 1.16, 1.18 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 263b (Certified Value 15.1 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.58 | 15.0968 | 16.6263 | 17.9 | 15.15 | 14 | 15.9 | 14.7 | 14.05 | 15.2917 | 15.04 | 15 | 12.3 | 15.41 | 14.8 | 12.8 | 14.6 |  |  |  | 1
| 1 | 2 | 14.66 | 14.9582 | 16.5357 | 17.9 | 15.3 | 14.1 | 16 | 14.8 | 14.05 | 15.2327 | 14.76 | 15.3 | 11.5 | 15.49 | 15.1 | 13.2 | 15.3 |  |  |  | 28
| 1 | 3 | 15.08 | 14.9608 | 16.4511 | 17.9 | 15.3 | 14.1 | 16.05 | 14.25 | 14.75 | 15.411 | 14.48 | 15.1 | 12.9 | 15.74 | 15 | 12.9 | 15 |  |  |  | 16
| 1 | 4 | 14.97 | 15.5147 | 16.6678 | 18 | 15.35 | 13.8 | 15.9 | 14.7 | 14.3 | 15.6885 | 14.55 | 14.9 | 11.9 | 15.6 | 14.8 | 13.2 | 14.6 |  |  |  | 15.0994
| 1 | 5 | 14.64 | 15.3255 | 16.632 | 17.6 | 14.75 | 13.5 | 16.5 | 14.5 | 14.35 | 15.6635 | 14.47 | 15.3 | 14.7 | 15.56 | 15.2 | 14.2 | 15.1 |  |  |  | 62
| 1 | 6 | 15.12 | 14.4704 | 16.58 | 18.2 | 15.95 | 13.85 | 16.05 | 14.95 | 14.55 | 15.5339 | 14.42 | 15.2 | 13.1 | 15.84 | 14.8 | 13.3 | 15.2
| Mean |  | 14.8417 | 15.0544 | 16.5822 | 17.9167 | 15.3 | 13.8917 | 16.0667 | 14.65 | 14.3417 | 15.4702 | 14.62 | 15.1333 | 12.7333 | 15.6067 | 14.95 | 13.2667 | 14.9667
| Median |  | 14.815 | 15.0288 | 16.6031 | 17.9 | 15.3 | 13.925 | 16.025 | 14.7 | 14.325 | 15.4724 | 14.515 | 15.15 | 12.6 | 15.58 | 14.9 | 13.2 | 15.05
| Std Dev. |  | 0.24203 | 0.35954 | 0.078875 | 0.19408 | 0.3873 | 0.22895 | 0.22286 | 0.24495 | 0.27644 | 0.19018 | 0.23791 | 0.1633 | 1.1343 | 0.15921 | 0.17607 | 0.49666 | 0.30111
| Rel.Std.Dev. |  | 1.631% | 2.388% | 0.4757% | 1.083% | 2.531% | 1.648% | 1.387% | 1.672% | 1.928% | 1.229% | 1.627% | 1.079% | 8.908% | 1.02% | 1.178% | 3.744% | 2.012%
| PDM3 |  | -1.707% | -0.2979% | 9.82% | 18.66% | 1.329% | -7.998% | 6.406% | -2.976% | -5.018% | 2.456% | -3.175% | 0.225% | -15.67% | 3.36% | -0.9892% | -12.14% | -0.8788%
| Z-Score (Absolute) |  | 0.17 | 0.12 | 2.16 | 3.95 | 0.45 | 1.44 | 1.47 | 0.42 | 0.84 | 0.67 | 0.46 | 0.22 | 2.99 | 0.86 | 0.02 | 2.28 | 0
| Analytical results for Ti in OREAS 263b (Certified Value 0.315 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.2865 | 0.30138 | 0.32394 | 0.333 | 0.25 | 0.312 | 0.319 | 0.341 | 0.324 | 0.315 | 0.30708 | 0.1806 | 0.227 | 0.303 | 0.316 | 0.286 | 0.336 | 0.34 | 0.34 | 0.29065 | 1
| 1 | 2 | 0.297 | 0.29666 | 0.32501 | 0.287 | 0.25 | 0.315 | 0.327 | 0.342 | 0.32 | 0.327 | 0.29857 | 0.1802 | 0.215 | 0.3 | 0.3308 | 0.293 | 0.331 | 0.34 | 0.37 | 0.30057 | 29
| 1 | 3 | 0.3023 | 0.29572 | 0.32759 | 0.324 | 0.25 | 0.316 | 0.322 | 0.343 | 0.314 | 0.32 | 0.27256 | 0.1792 | 0.214 | 0.318 | 0.336 | 0.291 | 0.298 | 0.34 | 0.375 | 0.29012 | 16
| 1 | 4 | 0.2888 | 0.29743 | 0.32491 | 0.317 | 0.25 | 0.312 | 0.326 | 0.346 | 0.325 | 0.321 | 0.30252 | 0.1778 | 0.21 | 0.298 | 0.3339 | 0.284 | 0.325 | 0.34 | 0.36 | 0.28794 | 0.31503
| 1 | 5 | 0.3279 | 0.2989 | 0.3241 | 0.313 | 0.24 | 0.313 | 0.32 | 0.335 | 0.334 | 0.319 | 0.27505 | 0.1778 | 0.213 | 0.339 | 0.3175 | 0.291 | 0.338 | 0.35 | 0.375 | 0.29996 | 63
| 1 | 6 | 0.3004 | 0.29829 | 0.32374 | 0.352 | 0.24 | 0.318 | 0.323 | 0.34 | 0.328 | 0.317 | 0.31583 | 0.1811 | 0.203 | 0.335 | 0.3322 | 0.295 | 0.32 | 0.34 | 0.38 | 0.30502
| Mean |  | 0.30048 | 0.29806 | 0.32488 | 0.321 | 0.24667 | 0.31433 | 0.32283 | 0.34117 | 0.32417 | 0.31983 | 0.29527 | 0.17945 | 0.21367 | 0.3155 | 0.32773 | 0.29 | 0.32467 | 0.34167 | 0.36667 | 0.29571
| Median |  | 0.2987 | 0.29786 | 0.3245 | 0.3205 | 0.25 | 0.314 | 0.3225 | 0.3415 | 0.3245 | 0.3195 | 0.30055 | 0.1797 | 0.2135 | 0.3105 | 0.3315 | 0.291 | 0.328 | 0.34 | 0.3725 | 0.2953
| Std Dev. |  | 0.014827 | 0.0019813 | 0.0014267 | 0.021679 | 0.005164 | 0.0024221 | 0.0031885 | 0.003656 | 0.006824 | 0.0041191 | 0.017609 | 0.0014223 | 0.0078401 | 0.018119 | 0.0086961 | 0.0041952 | 0.014692 | 0.0040825 | 0.01472 | 0.0070078
| Rel.Std.Dev. |  | 4.934% | 0.6647% | 0.4391% | 6.754% | 2.094% | 0.7706% | 0.9877% | 1.072% | 2.105% | 1.288% | 5.964% | 0.7926% | 3.669% | 5.743% | 2.653% | 1.447% | 4.525% | 1.195% | 4.014% | 2.37%
| PDM3 |  | -4.617% | -5.385% | 3.128% | 1.896% | -21.7% | -0.2203% | 2.478% | 8.297% | 2.901% | 1.526% | -6.271% | -43.04% | -32.18% | 0.15% | 4.033% | -7.944% | 3.06% | 8.456% | 16.39% | -6.133%
| Z-Score (Absolute) |  | 0.63 | 0.72 | 0.26 | 0.12 | 2.6 | 0.12 | 0.19 | 0.86 | 0.24 | 0.08 | 0.82 | 5.07 | 3.81 | 0.08 | 0.37 | 1.01 | 0.26 | 0.88 | 1.8 | 0.81
| Analytical results for Tl in OREAS 263b (Certified Value 1.16 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 1.18 | 1.2144 | 1.16 | 1.16 | 1.15 | 1.22 | 1.11 | 1.17 | 1.1523 | 1.119 | 1.15 | 1.07 | 1.2 | 1.13 | 1.07 | 1 | 2.52 |  |  |  | 1
| 1 | 2 | 1.18 | 1.2002 | 1.15 | 1.1 | 1.17 | 1.19 | 1.14 | 1.21 | 1.1555 | 1.111 | 1.16 | 1.04 | 1.2 | 1.11 | 1.1 | 1.1 | 1.612 |  |  |  | 30
| 1 | 3 | 1.19 | 1.1584 | 1.16 | 1.14 | 1.15 | 1.2 | 1.1 | 1.25 | 1.1758 | 1.089 | 1.15 | 1.06 | 1.23 | 1.12 | 1.07 | 1.1 | 1.299 |  |  |  | 16
| 1 | 4 | 1.17 | 1.1864 | 1.18 | 1.13 | 1.17 | 1.22 | 1.07 | 1.2 | 1.1998 | 1.099 | 1.15 | 1.16 | 1.24 | 1.1 | 1.09 | 1 | 1.036 |  |  |  | 1.1576
| 1 | 5 | 1.17 | 1.2275 | 1.15 | 1.12 | 1.12 | 1.2 | 1.1 | 1.21 | 1.1932 | 1.085 | 1.15 | 1.17 | 1.21 | 1.16 | 1.16 | 1.1 | 0.104 |  |  |  | 64
| 1 | 6 | 1.18 | 1.1994 | 1.16 | 1.13 | 1.14 | 1.22 | 1.12 | 1.22 | 1.1845 | 1.103 | 1.12 | 1.09 | 1.26 | 1.15 | 1.08 | 1.1
| Mean |  | 1.1783 | 1.1977 | 1.16 | 1.13 | 1.15 | 1.2083 | 1.1067 | 1.21 | 1.1769 | 1.101 | 1.1467 | 1.0983 | 1.2233 | 1.1283 | 1.095 | 1.0667 | 1.3142
| Median |  | 1.18 | 1.1998 | 1.16 | 1.13 | 1.15 | 1.21 | 1.105 | 1.21 | 1.1801 | 1.101 | 1.15 | 1.08 | 1.22 | 1.125 | 1.085 | 1.1 | 1.299
| Std Dev. |  | 0.0075277 | 0.023833 | 0.010954 | 0.02 | 0.018974 | 0.013292 | 0.023381 | 0.026077 | 0.019548 | 0.0129 | 0.013663 | 0.054191 | 0.024221 | 0.023166 | 0.033912 | 0.05164 | 0.87819
| Rel.Std.Dev. |  | 0.6388% | 1.99% | 0.9443% | 1.77% | 1.65% | 1.1% | 2.113% | 2.155% | 1.661% | 1.172% | 1.192% | 4.934% | 1.98% | 2.053% | 3.097% | 4.841% | 66.82%
| PDM3 |  | 1.787% | 3.463% | 0.2037% | -2.388% | -0.6601% | 4.379% | -4.403% | 4.523% | 1.66% | -4.893% | -0.9481% | -5.123% | 5.675% | -2.532% | -5.411% | -7.859% | 13.52%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.14 | 0.28 | 0 | 0.82 | 0.61 | 0.85 | 0.38 | 0.69 | 0.05 | 0.73 | 1.04 | 0.31 | 0.78 | 1.18 | 2.32
| Analytical results for Tm in OREAS 263b (Certified Value 0.24 ppm) | [21 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.23 | 0.23474 | 0.24849 | 0.3 | 0.2 | 0.25 | 0.2 | 0.22 | 0.25 | [11 empty] | 1
| 1 | 2 | 0.22 | 0.21272 | 0.25884 | 0.3 | 0.2 | 0.23 | 0.2 | 0.24 | 0.25 | [11 empty] | 6
| 1 | 3 | 0.25 | 0.2129 | 0.26784 | 0.3 | 0.2 | 0.25 | 0.2 | 0.22 | 0.25 | [11 empty] | 16
| 1 | 4 | 0.24 | 0.22754 | 0.26127 | 0.3 | 0.2 | 0.24 | 0.2 | 0.2 | 0.25 | [11 empty] | 0.23779
| 1 | 5 | 0.24 | 0.2317 | 0.25695 | 0.3 | 0.2 | 0.23 | 0.2 | 0.23 | 0.25 | [11 empty] | 65
| 1 | 6 | 0.23 | 0.23379 | 0.24363 | 0.3 | 0.2 | 0.23 | 0.2 | 0.22 | 0.25
| Mean |  | 0.235 | 0.22557 | 0.25617 | 0.3 | 0.2 | 0.23833 | 0.2 | 0.22167 | 0.25
| Median |  | 0.235 | 0.22962 | 0.25789 | 0.3 | 0.2 | 0.235 | 0.2 | 0.22 | 0.25
| Std Dev. |  | 0.010488 | 0.010187 | 0.0087883 | 0 | 3.0405e-17 | 0.0098319 | 3.0405e-17 | 0.013292 | 0
| Rel.Std.Dev. |  | 4.463% | 4.516% | 3.431% | 0% | 1.52e-14% | 4.125% | 1.52e-14% | 5.996% | 0%
| PDM3 |  | -1.173% | -5.14% | 7.73% | 26.16% | -15.89% | 0.2287% | -15.89% | -6.78% | 5.135%
| Z-Score (Absolute) |  | 0.09 | 0.61 | 1.08 | NA | NA | 0.09 | NA | 0.83 | 0.74
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 263b (Certified Value 2.49 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.43 | 2.6412 | 2.4013 | 2.6 | 2.4 | 2.4 | 2.6 | 2.5 | 2.4 | 2.4465 | 2.15 | 2.51 | 2.2 | 2.7 | 2.3 | 2.38 | 2.5 |  |  |  | 1
| 1 | 2 | 2.46 | 2.6237 | 2.367 | 2.6 | 2.4 | 2.4 | 2.6 | 2.5 | 2.6 | 2.3833 | 2.1 | 2.57 | 2.1 | 2.74 | 2.4 | 2.45 | 2.6 |  |  |  | 32
| 1 | 3 | 2.51 | 2.6338 | 2.4114 | 2.6 | 2.5 | 2.4 | 2.6 | 2.3 | 2.7 | 2.4371 | 2.1 | 2.47 | 2.3 | 2.82 | 2.3 | 2.37 | 2.5 |  |  |  | 16
| 1 | 4 | 2.46 | 2.6369 | 2.4654 | 2.6 | 2.4 | 2.3 | 2.6 | 2.4 | 2.5 | 2.5488 | 2.08 | 2.54 | 2.2 | 2.76 | 2.3 | 2.43 | 2.6 |  |  |  | 2.4863
| 1 | 5 | 2.47 | 2.5225 | 2.4229 | 2.6 | 2.4 | 2.2 | 2.6 | 2.3 | 2.6 | 2.443 | 2.09 | 2.49 | 2.5 | 2.67 | 2.4 | 2.62 | 2.6 |  |  |  | 66
| 1 | 6 | 2.53 | 2.5164 | 2.4114 | 2.6 | 2.5 | 2.3 | 2.6 | 2.5 | 2.6 | 2.5623 | 2.08 | 2.53 | 2.4 | 2.78 | 2.4 | 2.43 | 2.6
| Mean |  | 2.4767 | 2.5957 | 2.4132 | 2.6 | 2.4333 | 2.3333 | 2.6 | 2.4167 | 2.5667 | 2.4702 | 2.1 | 2.5183 | 2.2833 | 2.745 | 2.35 | 2.4467 | 2.5667
| Median |  | 2.465 | 2.6287 | 2.4114 | 2.6 | 2.4 | 2.35 | 2.6 | 2.45 | 2.6 | 2.4447 | 2.095 | 2.52 | 2.25 | 2.75 | 2.35 | 2.43 | 2.6
| Std Dev. |  | 0.036697 | 0.059425 | 0.031931 | 0 | 0.05164 | 0.08165 | 0 | 0.098319 | 0.10328 | 0.070163 | 0.026077 | 0.036009 | 0.1472 | 0.054314 | 0.054772 | 0.09048 | 0.05164
| Rel.Std.Dev. |  | 1.482% | 2.289% | 1.323% | 0% | 2.122% | 3.499% | 0% | 4.068% | 4.024% | 2.84% | 1.242% | 1.43% | 6.447% | 1.979% | 2.331% | 3.698% | 2.012%
| PDM3 |  | -0.3882% | 4.401% | -2.94% | 4.572% | -2.131% | -6.153% | 4.572% | -2.801% | 3.232% | -0.6503% | -15.54% | 1.288% | -8.164% | 10.4% | -5.483% | -1.595% | 3.232%
| Z-Score (Absolute) |  | 0.05 | 0.88 | 0.4 | 0.91 | 0.26 | 0.96 | 0.91 | 0.37 | 0.67 | 0 | 2.59 | 0.34 | 1.31 | 1.92 | 0.84 | 0.16 | 0.67
| Analytical results for V in OREAS 263b (Certified Value 76 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 75 | 77.8815 | 70.5495 | 34 | 75 | 74 | 76 | 79 | 76 | 75 | 78.3074 | 77 | 77 | 78 | 69 | 78 | 63 | 74.15 | 75 | 74.62 | 1
| 1 | 2 | 74 | 76.9976 | 74.4765 | 48 | 77 | 74 | 77 | 79 | 75 | 75 | 77.7338 | 77 | 76 | 79 | 76 | 78 | 66 | 74.41 | 75 | 76.241 | 33
| 1 | 3 | 75 | 75.7562 | 74.0145 | 46 | 77 | 74 | 75 | 79 | 76 | 74 | 77.4264 | 76 | 75 | 81 | 73 | 78 | 62 | 73.76 | 75 | 74.864 | 16
| 1 | 4 | 75 | 77.3146 | 71.6415 | 51 | 75 | 75 | 76 | 80 | 75 | 77 | 77.2518 | 76 | 74 | 82 | 76 | 77 | 61 | 74.95 | 75 | 75.385 | 75.8243
| 1 | 5 | 74 | 77.5671 | 77.595 | 55 | 73 | 74 | 78 | 79 | 77 | 74 | 78.3336 | 77 | 76 | 84 | 79 | 79 | 70 | 73.54 | 75 | 76.858 | 67
| 1 | 6 | 70 | 77.0605 | 65.0475 | 45 | 72 | 75 | 75 | 78 | 76 | 77 | 77.5175 | 76 | 75 | 84 | 72 | 79 | 62 | 74.72 | 75 | 76.7
| Mean |  | 73.8333 | 77.0963 | 72.2208 | 46.5 | 74.8333 | 74.3333 | 76.1667 | 79 | 75.8333 | 75.3333 | 77.7617 | 76.5 | 75.5 | 81.3333 | 74.1667 | 78.1667 | 64 | 74.255 | 75 | 75.778
| Median |  | 74.5 | 77.1876 | 72.828 | 47 | 75 | 74 | 76 | 79 | 76 | 75 | 77.6256 | 76.5 | 75.5 | 81.5 | 74.5 | 78 | 62.5 | 74.28 | 75 | 75.813
| Std Dev. |  | 1.9408 | 0.73386 | 4.2842 | 7.1204 | 2.0412 | 0.5164 | 1.169 | 0.63246 | 0.75277 | 1.3663 | 0.45992 | 0.54772 | 1.0488 | 2.5033 | 3.5449 | 0.75277 | 3.4059 | 0.54599 | 0 | 0.95516
| Rel.Std.Dev. |  | 2.629% | 0.9519% | 5.932% | 15.31% | 2.728% | 0.6947% | 1.535% | 0.8006% | 0.9927% | 1.814% | 0.5914% | 0.716% | 1.389% | 3.078% | 4.78% | 0.963% | 5.322% | 0.7353% | 0% | 1.26%
| PDM3 |  | -2.626% | 1.678% | -4.752% | -38.67% | -1.307% | -1.966% | 0.4515% | 4.188% | 0.01192% | -0.6475% | 2.555% | 0.8911% | -0.4277% | 7.266% | -2.186% | 3.089% | -15.59% | -2.07% | -1.087% | -0.06106%
| Z-Score (Absolute) |  | 0.89 | 0.94 | 1.79 | 16.17 | 0.33 | 0.61 | 0.42 | 2 | 0.23 | 0.05 | 1.31 | 0.61 | 0.05 | 3.31 | 0.7 | 1.54 | 6.38 | 0.65 | 0.23 | 0.2
| Analytical results for W in OREAS 263b (Certified Value 42.5 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 41.1 | 46.6242 | 41.9 | 29.8 | 44.3 | 42.7 | 45.3 | 43.8 | 43.9 | 27.2461 | 36.42 | 40 | 39.9 | 46.4 | 40.1 | 44.1 | 34.29 |  |  |  | 1
| 1 | 2 | 40.8 | 46.3591 | 41.4 | 30.9 | 45.1 | 44.9 | 46.6 | 42.4 | 44.6 | 31.2298 | 36.32 | 37.5 | 39.9 | 43.8 | 43.6 | 48.9 | 36.02 |  |  |  | 34
| 1 | 3 | 42.8 | 44.6249 | 43.4 | 30.3 | 45.6 | 43.2 | 44.9 | 41.4 | 46.4 | 28.7664 | 36.93 | 38.3 | 40.7 | 47.1 | 42.8 | 44.3 | 35.11 |  |  |  | 16
| 1 | 4 | 44.1 | 44.0121 | 42 | 29.8 | 45.5 | 43.8 | 47.5 | 41.5 | 45 | 33.8979 | 36.72 | 37.2 | 41.9 | 45.8 | 42 | 44 | 35 |  |  |  | 42.5159
| 1 | 5 | 42.2 | 45.986 | 43.2 | 29.5 | 45.4 | 38.5 | 47.7 | 45.1 | 45.8 | 28.9764 | 34.63 | 36.2 | 41.6 | 46.3 | 41.1 | 47.5 | 37.74 |  |  |  | 68
| 1 | 6 | 43.3 | 44.5769 | 42.8 | 27.8 | 46.4 | 44.3 | 47.8 | 44.5 | 44.7 | 33.3918 | 35.32 | 39 | 41.7 | 45.9 | 39 | 45 | 35.35
| Mean |  | 42.3833 | 45.3638 | 42.45 | 29.6833 | 45.3833 | 42.9 | 46.6333 | 43.1167 | 45.0667 | 30.5847 | 36.0567 | 38.0333 | 40.95 | 45.8833 | 41.4333 | 45.6333 | 35.585
| Median |  | 42.5 | 45.3055 | 42.4 | 29.8 | 45.45 | 43.5 | 47.05 | 43.1 | 44.85 | 30.1031 | 36.37 | 37.9 | 41.15 | 46.1 | 41.55 | 44.65 | 35.23
| Std Dev. |  | 1.2766 | 1.0917 | 0.79937 | 1.0458 | 0.68532 | 2.2917 | 1.2675 | 1.5741 | 0.89815 | 2.6949 | 0.89243 | 1.3574 | 0.91159 | 1.1197 | 1.7143 | 2.0666 | 1.1944
| Rel.Std.Dev. |  | 3.012% | 2.406% | 1.883% | 3.523% | 1.51% | 5.342% | 2.718% | 3.651% | 1.993% | 8.811% | 2.475% | 3.569% | 2.226% | 2.44% | 4.137% | 4.529% | 3.357%
| PDM3 |  | -0.3117% | 6.699% | -0.1549% | -30.18% | 6.744% | 0.9035% | 9.685% | 1.413% | 6% | -28.06% | -15.19% | -10.54% | -3.683% | 7.921% | -2.546% | 7.333% | -16.3%
| Z-Score (Absolute) |  | 0.02 | 0.67 | 0 | 2.93 | 0.67 | 0.1 | 0.96 | 0.15 | 0.6 | 2.73 | 1.47 | 1.02 | 0.34 | 0.79 | 0.23 | 0.73 | 1.58
| Analytical results for Y in OREAS 263b (Certified Value 14.3 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  | 1
1
| 1 | 1 | 14.08 | 14.0546 | 15.4998 | 18 | 11.5 | 15.1 | 15.1 | 13.9 | 14.5 | 14.1 | 15.6433 | 11.79 | 13.9 | 11.8 | 15.21 | 12.6 | 15.5 | 15.4 | 13.673 |  | 1
| 1 | 2 | 14.28 | 14.034 | 15.5702 | 17.1 | 11.8 | 16 | 15 | 14.6 | 14.4 | 14.8 | 15.469 | 11.72 | 14.3 | 11.3 | 16.09 | 13.2 | 15.4 | 15.5 | 12.896 |  | 35
| 1 | 3 | 14.44 | 13.8019 | 15.4064 | 17.6 | 11.9 | 15.3 | 14.9 | 14.2 | 13.8 | 14.8 | 15.2688 | 11.41 | 13.9 | 12.2 | 16.03 | 12.6 | 16.1 | 15.3 | 12.136 |  | 16
| 1 | 4 | 14.34 | 14.2984 | 15.455 | 17.3 | 12.7 | 15 | 14.7 | 14.6 | 14 | 14.4 | 15.5295 | 11.59 | 13.8 | 11.3 | 16.08 | 13 | 15.5 | 15.4 | 12.473 |  | 14.3174
| 1 | 5 | 14.35 | 14.0915 | 15.456 | 16.7 | 11.9 | 14.9 | 13.8 | 14.4 | 14.1 | 14.4 | 15.2777 | 11.21 | 14 | 13.7 | 15.74 | 13.1 | 15.3 | 15.5 | 13.878 |  | 69
| 1 | 6 | 14.21 | 13.6257 | 15.4643 | 18.1 | 11.2 | 15.1 | 15.1 | 14.1 | 14.6 | 14.2 | 15.6681 | 11.48 | 13.9 | 12.4 | 16.9 | 13.1 | 15.2 | 15.3 | 13.843
| Mean |  | 14.2833 | 13.9844 | 15.4753 | 17.4667 | 11.8333 | 15.2333 | 14.7667 | 14.3 | 14.2333 | 14.45 | 15.4761 | 11.5333 | 13.9667 | 12.1167 | 16.0083 | 12.9333 | 15.5 | 15.4 | 13.1498
| Median |  | 14.31 | 14.0443 | 15.4601 | 17.45 | 11.85 | 15.1 | 14.95 | 14.3 | 14.25 | 14.4 | 15.4992 | 11.535 | 13.9 | 12 | 16.055 | 13.05 | 15.45 | 15.4 | 13.2845
| Std Dev. |  | 0.12565 | 0.2364 | 0.055243 | 0.53914 | 0.50465 | 0.39833 | 0.49666 | 0.28284 | 0.31411 | 0.29496 | 0.17324 | 0.21285 | 0.17512 | 0.89759 | 0.55076 | 0.26583 | 0.31623 | 0.089443 | 0.75297
| Rel.Std.Dev. |  | 0.8797% | 1.69% | 0.357% | 3.087% | 4.265% | 2.615% | 3.363% | 1.978% | 2.207% | 2.041% | 1.119% | 1.846% | 1.254% | 7.408% | 3.44% | 2.055% | 2.04% | 0.5808% | 5.726%
| PDM3 |  | -0.2379% | -2.326% | 8.087% | 22% | -17.35% | 6.397% | 3.138% | -0.1215% | -0.5871% | 0.9262% | 8.093% | -19.45% | -2.45% | -15.37% | 11.81% | -9.667% | 8.26% | 7.561% | -8.155%
| Z-Score (Absolute) |  | 0.01 | 0.19 | 0.69 | 1.86 | 1.45 | 0.55 | 0.27 | 0 | 0.04 | 0.09 | 0.69 | 1.62 | 0.2 | 1.28 | 1 | 0.8 | 0.7 | 0.64 | 0.67
| Analytical results for Yb in OREAS 263b (Certified Value 1.63 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 1.61 | 1.6321 | 1.2872 | 2.1 | 1.3 | 1.67 | 1.4 | 1.55 | 1.95 | [11 empty] | 1
| 1 | 2 | 1.63 | 1.3921 | 1.3363 | 2 | 1.4 | 1.75 | 1.4 | 1.57 | 1.95 | [11 empty] | 7
| 1 | 3 | 1.67 | 1.5613 | 1.3684 | 2.1 | 1.4 | 1.71 | 1.4 | 1.49 | 1.9 | [11 empty] | 16
| 1 | 4 | 1.65 | 1.6365 | 1.3316 | 2 | 1.3 | 1.66 | 1.4 | 1.56 | 1.95 | [11 empty] | 1.6273
| 1 | 5 | 1.56 | 1.5986 | 1.3412 | 2 | 1.6 | 1.77 | 1.4 | 1.59 | 1.95 | [11 empty] | 70
| 1 | 6 | 1.63 | 1.5613 | 1.3608 | 2.1 | 1.5 | 1.8 | 1.4 | 1.59 | 1.9
| Mean |  | 1.625 | 1.5637 | 1.3376 | 2.05 | 1.4167 | 1.7267 | 1.4 | 1.5583 | 1.9333
| Median |  | 1.63 | 1.58 | 1.3387 | 2.05 | 1.4 | 1.73 | 1.4 | 1.565 | 1.95
| Std Dev. |  | 0.037815 | 0.090187 | 0.028557 | 0.054772 | 0.1169 | 0.056095 | 2.4324e-16 | 0.037103 | 0.02582
| Rel.Std.Dev. |  | 2.327% | 5.768% | 2.135% | 2.672% | 8.252% | 3.249% | 1.737e-14% | 2.381% | 1.336%
| PDM3 |  | -0.1403% | -3.91% | -17.8% | 25.98% | -12.94% | 6.107% | -13.97% | -4.237% | 18.81%
| Z-Score (Absolute) |  | 0.25 | 0 | 0.94 | 2.01 | 0.61 | 0.67 | 0.68 | 0.02 | 1.53
| Analytical results for Zn in OREAS 263b (Certified Value 44 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 43 | 44.2016 | 44.784 | 49 | 45 | 42 | 43 | 45 | 40 | 20 | 43.175 | 43 | 45 | 43.2 | 44 | 41.8 | 34.6 | 71.97 | 46 | 42.776 | 1
| 1 | 2 | 44 | 42.7259 | 42.96 | 48 | 42 | 42 | 44 | 46 | 41 | 38 | 42.993 | 43 | 43 | 41.8 | 45 | 41.3 | 38.8 | 68.23 | 46 | 43.724 | 15
| 1 | 3 | 44 | 43.2659 | 45.24 | 49 | 43 | 42 | 43 | 46 | 41 | 37 | 41.7815 | 43 | 43 | 48.2 | 46 | 41.5 | 35.5 | 70.06 | 48 | 43.577 | 16
| 1 | 4 | 45 | 44.3794 | 42.336 | 47 | 42 | 43 | 44 | 46 | 41 | 26 | 46.036 | 43 | 43 | 44.1 | 45 | 41.6 | 38.3 | 69.41 | 44 | 43.502 | 44.0285
| 1 | 5 | 44 | 44.5744 | 45.852 | 49 | 40 | 42 | 44 | 46 | 45 | 37 | 46.0267 | 43 | 44 | 48.5 | 46 | 42.8 | 37.2 | 69.05 | 46 | 44.723 | 71
| 1 | 6 | 44 | 43.7376 | 44.172 | 50 | 41 | 43 | 44 | 45 | 41 | 22 | 43.9091 | 44 | 43 | 48.9 | 45 | 42.6 | 35.7 | 67.07 | 46 | 43.355
| Mean |  | 44 | 43.8142 | 44.224 | 48.6667 | 42.1667 | 42.3333 | 43.6667 | 45.6667 | 41.5 | 30 | 43.9869 | 43.1667 | 43.5 | 45.7833 | 45.1667 | 41.9333 | 36.6833 | 69.2983 | 46 | 43.6095
| Median |  | 44 | 43.9696 | 44.478 | 49 | 42 | 42 | 44 | 46 | 41 | 31.5 | 43.542 | 43 | 43 | 46.15 | 45 | 41.7 | 36.45 | 69.23 | 46 | 43.5395
| Std Dev. |  | 0.63246 | 0.71252 | 1.3536 | 1.0328 | 1.7224 | 0.5164 | 0.5164 | 0.5164 | 1.7607 | 8.2704 | 1.7249 | 0.40825 | 0.83666 | 3.1083 | 0.75277 | 0.6186 | 1.6774 | 1.6673 | 1.2649 | 0.63641
| Rel.Std.Dev. |  | 1.437% | 1.626% | 3.061% | 2.122% | 4.085% | 1.22% | 1.183% | 1.131% | 4.243% | 27.57% | 3.921% | 0.9457% | 1.923% | 6.789% | 1.667% | 1.475% | 4.573% | 2.406% | 2.75% | 1.459%
| PDM3 |  | -0.06473% | -0.4868% | 0.444% | 10.53% | -4.229% | -3.85% | -0.8218% | 3.721% | -5.743% | -31.86% | -0.09456% | -1.957% | -1.2% | 3.986% | 2.585% | -4.759% | -16.68% | 57.39% | 4.478% | -0.9517%
| Z-Score (Absolute) |  | 0.12 | 0.03 | 0.22 | 2.21 | 0.71 | 0.63 | 0.03 | 0.87 | 1.01 | 6.18 | 0.11 | 0.26 | 0.11 | 0.92 | 0.64 | 0.81 | 3.17 | 11.49 | 1.02 | 0.06
| Analytical results for Zr in OREAS 263b (Certified Value 110 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 110.2 | 112.0363 | 112.8935 | 24 | 103 | 119.5 | 121 | 114.5 | 121 | 104.5 | 109.811 | 92.5 | 103 | 97.9 | 117.6 | 101 | 122 | 104.2 | 110 |  | 1
| 1 | 2 | 112.8 | 108.4978 | 116.4215 | 87 | 103 | 121.5 | 123 | 116 | 120 | 103 | 106.5814 | 91.7 | 100 | 93.1 | 120.4 | 102.5 | 114 | 97.5 | 110 |  | 16
| 1 | 3 | 117.5 | 111.6879 | 119.6345 | 66 | 108 | 123 | 119.5 | 113.5 | 113 | 114 | 104.5843 | 88.5 | 104 | 112.8 | 123.3 | 101 | 122 | 104.5 | 110 |  | 16
| 1 | 4 | 110.6 | 108.0462 | 118.595 | 92 | 102 | 119.5 | 121.5 | 117 | 115.5 | 105 | 111.8181 | 89.4 | 100 | 114 | 122 | 100.6 | 112 | 97.8 | 105 |  | 110.3413
| 1 | 5 | 114.3 | 109.6148 | 119.5925 | 103 | 99 | 123 | 113 | 113 | 115.5 | 104.5 | 106.7835 | 89.2 | 103 | 119.9 | 119.4 | 103.1 | 116 | 119.5 | 110 |  | 72
| 1 | 6 | 115.5 | 108.6203 | 115.6655 | 22 | 99 | 122 | 124 | 114.5 | 119.5 | 104.5 | 113.7738 | 90.3 | 108 | 117.1 | 121.6 | 103.6 | 114 | 95.2 | 110
| Mean |  | 113.4833 | 109.7506 | 117.1338 | 65.6667 | 102.3333 | 121.4167 | 120.3333 | 114.75 | 117.4167 | 105.9167 | 108.892 | 90.2667 | 103 | 109.1333 | 120.7167 | 101.9667 | 116.6667 | 103.1167 | 109.1667
| Median |  | 113.55 | 109.1176 | 117.5083 | 76.5 | 102.5 | 121.75 | 121.25 | 114.5 | 117.5 | 104.5 | 108.2972 | 89.85 | 103 | 113.4 | 121 | 101.75 | 115 | 101 | 110
| Std Dev. |  | 2.8435 | 1.7174 | 2.6487 | 35.172 | 3.3267 | 1.5943 | 3.92 | 1.5083 | 3.1846 | 4.0177 | 3.5094 | 1.5526 | 2.9665 | 10.9524 | 2.0322 | 1.2628 | 4.3205 | 8.8768 | 2.0412
| Rel.Std.Dev. |  | 2.506% | 1.565% | 2.261% | 53.56% | 3.251% | 1.313% | 3.258% | 1.314% | 2.712% | 3.793% | 3.223% | 1.72% | 2.88% | 10.04% | 1.683% | 1.238% | 3.703% | 8.609% | 1.87%
| PDM3 |  | 2.848% | -0.5353% | 6.156% | -40.49% | -7.257% | 10.04% | 9.056% | 3.996% | 6.412% | -4.01% | -1.313% | -18.19% | -6.653% | -1.095% | 9.403% | -7.59% | 5.733% | -6.548% | -1.065%
| Z-Score (Absolute) |  | 0.43 | 0.06 | 0.79 | 4.29 | 0.67 | 1.21 | 1.1 | 0.55 | 0.81 | 0.32 | 0.03 | 1.87 | 0.61 | 0 | 1.14 | 0.71 | 0.74 | 0.6 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C686: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C813: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C887: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C905: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C923: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C978: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C996: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1014: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1033: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1051: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1069: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1087: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1105: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1123: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 263b (Certified Value 0.275 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.26 | 0.27292 | 0.072327 | 0.301 | < 0.5 | 0.26 | 0.28 | 0.25 | 0.28 | 0.26 | 0.30208 | < 0.3 | 0.16 | 0.25 | 0.257 | 0.26 | < 0.5 | 0.3 | 0.11 |  |  | 1
| 1 | 2 | 0.26 | 0.25727 | 0.12409 | 0.306 | < 0.5 | 0.26 | 0.29 | 0.25 | 0.28 | 0.28 | 0.32906 | < 0.3 | 0.13 | 0.254 | 0.262 | 0.27 | < 0.5 | 0.3 | 0.007 |  |  | 22
| 1 | 3 | 0.27 | 0.274 | 0.094507 | 0.292 | < 0.5 | 0.27 | 0.29 | 0.25 | 0.28 | 0.28 | 0.3166 | < 0.3 | 0.13 | 0.241 | 0.27 | 0.26 | < 0.5 | 0.3 | 0.139 |  |  | 16
| 1 | 4 | 0.27 | 0.27581 | 0.092 | 0.315 | < 0.5 | 0.27 | 0.29 | 0.25 | 0.28 | 0.26 | 0.33422 | < 0.3 | 0.15 | 0.256 | 0.281 | 0.28 | < 0.5 | 0.3 | 0.093 |  |  | 0.27462
| 1 | 5 | 0.26 | 0.25701 | 0.097875 | 0.314 | < 0.5 | 0.27 | 0.28 | 0.25 | 0.26 | 0.27 | 0.31702 | < 0.3 | 0.14 | 0.236 | 0.27 | 0.28 | < 0.5 | 0.3 | 0.062 |  |  | 74
| 1 | 6 | 0.26 | 0.28263 | 0.11036 | 0.292 | < 0.5 | 0.27 | 0.29 | 0.26 | 0.28 | 0.28 | 0.29673 | < 0.3 | 0.14 | 0.272 | 0.278 | 0.26 | < 0.5 | 0.3 | 0.119
| Mean |  | 0.26333 | 0.26994 | 0.098525 | 0.30333 |  | 0.26667 | 0.28667 | 0.25167 | 0.27667 | 0.27167 | 0.31595 |  | 0.14167 | 0.2515 | 0.26967 | 0.26833 |  | 0.3 | 0.088333
| Median |  | 0.26 | 0.27346 | 0.096191 | 0.3035 |  | 0.27 | 0.29 | 0.25 | 0.28 | 0.275 | 0.31681 |  | 0.14 | 0.252 | 0.27 | 0.265 |  | 0.3 | 0.1015
| Std Dev. |  | 0.005164 | 0.010474 | 0.017543 | 0.010191 |  | 0.005164 | 0.005164 | 0.0040825 | 0.008165 | 0.0098319 | 0.014622 |  | 0.01169 | 0.012645 | 0.009136 | 0.0098319 |  | 0 | 0.047555
| Rel.Std.Dev. |  | 1.961% | 3.88% | 17.81% | 3.36% |  | 1.936% | 1.801% | 1.622% | 2.951% | 3.619% | 4.628% |  | 8.252% | 5.028% | 3.388% | 3.664% |  | 0% | 53.84%
| PDM3 |  | -4.108% | -1.703% | -64.12% | 10.46% |  | -2.895% | 4.388% | -8.357% | 0.7468% | -1.074% | 15.05% |  | -48.41% | -8.418% | -1.802% | -2.288% |  | 9.243% | -67.83%
| Z-Score (Absolute) |  | 0.08 | 0.23 | 7.88 | 1.81 | 0.71 | 0.08 | 1.03 | 0.63 | 0.55 | 0.32 | 2.41 | 5.44 | 5.84 | 0.64 | 0.22 | 0.16 | 0.71 | NA | 8.36
Results from laboratory 1.28 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Al in OREAS 263b (Certified Value 0.967 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 0.9827 | 0.9968 | 1.24 | 1.02 | 0.95 | 1.06 | 0.94 | 0.91 | 0.92 | 1.0694 | 1.43 | 0.97 | 1.4675 | 0.82 | 1.84 | 0.88 | 1.03 | 1.0588 |  |  |  | 1
| 1 | 2 | 1.0022 | 0.99965 | 1.3 | 1.02 | 0.95 | 1.05 | 0.93 | 0.94 | 0.92 | 1.0664 | 1.41 | 0.96 | 1.4561 | 0.89 | 1.69 | 0.89 | 1.02 | 1.0569 |  |  |  | #N/A
| 1 | 3 | 1.0227 | 0.99478 | 1.24 | 1 | 0.9 | 1.06 | 0.95 | 0.89 | 0.94 | 0.99166 | 1.51 | 0.91 | 1.5156 | 0.89 | 1.73 | 0.9 | 1.04 | 0.98275 |  |  |  | 16
| 1 | 4 | 0.9942 | 0.99835 | 1.29 | 1.01 | 0.99 | 1.04 | 0.91 | 0.92 | 0.87 | 0.99273 | 1.48 | 0.89 | 1.4971 | 0.91 | 1.76 | 0.88 | 1.06 | 0.95326 |  |  |  | 0.96715
| 1 | 5 | 0.992 | 0.99437 | 1.31 | 1.01 | 0.92 | 1.04 | 0.93 | 0.94 | 0.93 | 1.0351 | 1.5 | 0.94 | 1.4784 | 0.87 | 1.56 | 0.89 | 1.01 | 0.95617 |  |  |  | 75
| 1 | 6 | 0.9914 | 1.0069 | 1.12 | 1.01 | 0.93 | 1.04 | 0.93 | 0.92 | 0.87 | 1.0482 | 1.49 | 0.98 | 1.4096 | 0.92 | 1.96 | 0.89 | 1.03 | 1.023
| Mean |  | 0.99753 | 0.99847 | 1.25 | 1.0117 | 0.94 | 1.0483 | 0.93167 | 0.92 | 0.90833 | 1.0339 | 1.47 | 0.94167 | 1.4707 | 0.88333 | 1.7567 | 0.88833 | 1.0317 | 1.0052
| Median |  | 0.9931 | 0.99758 | 1.265 | 1.01 | 0.94 | 1.045 | 0.93 | 0.92 | 0.92 | 1.0416 | 1.485 | 0.95 | 1.4729 | 0.89 | 1.745 | 0.89 | 1.03 | 1.0029
| Std Dev. |  | 0.013817 | 0.0045889 | 0.070427 | 0.0075277 | 0.030984 | 0.0098319 | 0.013292 | 0.018974 | 0.030605 | 0.03466 | 0.040497 | 0.035449 | 0.036685 | 0.03559 | 0.13574 | 0.0075277 | 0.017224 | 0.047894
| Rel.Std.Dev. |  | 1.385% | 0.4596% | 5.634% | 0.7441% | 3.296% | 0.9379% | 1.427% | 2.062% | 3.369% | 3.352% | 2.755% | 3.765% | 2.494% | 4.029% | 7.727% | 0.8474% | 1.67% | 4.765%
| PDM3 |  | 3.142% | 3.239% | 29.25% | 4.603% | -2.807% | 8.394% | -3.669% | -4.875% | -6.081% | 6.904% | 51.99% | -2.635% | 52.07% | -8.666% | 81.63% | -8.149% | 6.671% | 3.93%
| Z-Score (Absolute) |  | 0.04 | 0.03 | 2.54 | 0.1 | 0.63 | 0.48 | 0.72 | 0.84 | 0.96 | 0.33 | 4.78 | 0.61 | 4.79 | 1.21 | 7.71 | 1.16 | 0.31 | 0.03
| Analytical results for As in OREAS 263b (Certified Value 231 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 0
| 1 | 1 | 221 | 230.832 | 226.573 | 234.03 | 236 | 237 | 236 | 244 | 229 | 240 | 228 | 234.5623 | 219.4 | 225 | 231.3 | 240 | 208.8 | 222 | 216.6 | 230 | 245.312 | 1
| 1 | 2 | 226 | 226.2549 | 235.13 | 238.103 | 232 | 232 | 235 | 243 | 228 | 244 | 224 | 229.4723 | 215.7 | 223 | 234.4 | 241 | 220.6 | 227 | 199 | 230 | 246.749 | 10
| 1 | 3 | 228 | 232.5318 | 232.792 | 236.197 | 227 | 234 | 229 | 244 | 232 | 240 | 230 | 221.4945 | 223.8 | 234 | 230.7 | 241 | 219.2 | 219 | 198.7 | 226 | 238.93 | 16
| 1 | 4 | 223 | 231.8336 | 217 | 230.081 | 235 | 236 | 243 | 241 | 223 | 243 | 217 | 223.4508 | 219.2 | 228 | 224.7 | 241 | 219.5 | 227 | 209.9 | 228 | 237.644 | 231.2736
| 1 | 5 | 227 | 237.3983 | 234.536 | 240.465 | 235 | 235 | 234 | 241 | 224 | 243 | 229 | 232.3359 | 220.5 | 228 | 226.3 | 238 | 216.7 | 240 | 206.9 | 233 | 236.504 | 76
| 1 | 6 | 225 | 236.1607 | 230.192 | 237.099 | 226 | 235 | 236 | 241 | 229 | 243 | 216 | 228.7584 | 223.4 | 225 | 225.1 | 233 | 222 | 229 | 207 | 231 | 238.259
| Mean |  | 225 | 232.5019 | 229.3705 | 235.9958 | 231.8333 | 234.8333 | 235.5 | 242.3333 | 227.5 | 242.1667 | 224 | 228.3457 | 220.3333 | 227.1667 | 228.75 | 239 | 217.8 | 227.3333 | 206.35 | 229.6667 | 240.5663
| Median |  | 225.5 | 232.1827 | 231.492 | 236.648 | 233.5 | 235 | 235.5 | 242 | 228.5 | 243 | 226 | 229.1153 | 219.95 | 226.5 | 228.5 | 240.5 | 219.35 | 227 | 206.95 | 230 | 238.5945
| Std Dev. |  | 2.6077 | 3.9906 | 6.8278 | 3.5922 | 4.3551 | 1.7224 | 4.5056 | 1.5055 | 3.3912 | 1.7224 | 6.1644 | 5.038 | 3.0024 | 3.8687 | 3.9485 | 3.1623 | 4.7443 | 7.2296 | 6.7961 | 2.4221 | 4.3311
| Rel.Std.Dev. |  | 1.159% | 1.716% | 2.977% | 1.522% | 1.879% | 0.7335% | 1.913% | 0.6213% | 1.491% | 0.7112% | 2.752% | 2.206% | 1.363% | 1.703% | 1.726% | 1.323% | 2.178% | 3.18% | 3.293% | 1.055% | 1.8%
| PDM3 |  | -2.713% | 0.5311% | -0.8229% | 2.042% | 0.242% | 1.539% | 1.827% | 4.782% | -1.632% | 4.71% | -3.145% | -1.266% | -4.73% | -1.776% | -1.091% | 3.341% | -5.826% | -1.704% | -10.78% | -0.6948% | 4.018%
| Z-Score (Absolute) |  | 0.55 | 0.39 | 0 | 0.83 | 0.31 | 0.69 | 0.77 | 1.63 | 0.23 | 1.61 | 0.67 | 0.13 | 1.13 | 0.28 | 0.08 | 1.21 | 1.45 | 0.26 | 2.89 | 0.04 | 1.41
| Analytical results for B in OREAS 263b (Certified Value < 10 ppm) | [22 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 1
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 15g | [9 empty] | 1
| 1 | 1 | 5.936 | 5 | < 10 | 10 | 10 | 10 | 10 | < 10 | < 10 | < 20 | 10 | 4 | [9 empty] | 1
| 1 | 2 | 5.7338 | 5 | < 10 | 10 | 10 | 10 | < 10 | < 10 | < 10 | < 20 | 11 | 4 | [9 empty] | 1
| 1 | 3 | 5 | 6 | < 10 | 10 | 10 | 10 | < 10 | 10 | < 10 | < 20 | 11 | 4 | [9 empty] | 16
| 1 | 4 | 4.584 | 6 | < 10 | 10 | 10 | < 10 | 10 | < 10 | < 10 | < 20 |  | 4 | [9 empty] | < 10
| 1 | 5 | 5.8184 | 5 | < 10 | 10 | 10 | 10 | 10 | < 10 | < 10 | < 20 |  | 4 | [9 empty] | 77
| 1 | 6 | 5.3754 | 5 | < 10 | 10 | 10 | 10 | 10 | < 10 | < 10 | < 20 |  | 4
| Mean |  | 5.4079 | 5.3333 |  | 10 | 10 | 10 | 10 | 10 |  |  | 10.6667 | 4
| Median |  | 5.5546 | 5 |  | 10 | 10 | 10 | 10 | 10 |  |  | 11 | 4
| Std Dev. |  | 0.52911 | 0.5164 |  | 0 | 0 | 0 | 0 |  |  |  | 0.57735 | 0
| Rel.Std.Dev. |  | 9.784% | 9.682% |  | 0% | 0% | 0% | 0% |  |  |  | 5.413% | 0%
PDM3
| Z-Score (Absolute) |  | 0.54 | 0.57 | 0.67 | 0.94 | 0.94 | 0.67 | 0.4 | 0.4 | 0.67 | 0.94 | 1.16 | 1
| Analytical results for Ba in OREAS 263b (Certified Value 68 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 0.5g |  |  |  | 0
| 1 | 1 | 65 | 65.178 | 59.137 | 90 | 70 | 60 | 70 | 60 | 70 | 60 | 73.4938 | 87 | 70.8 | 101 | 54.2 | 108 | 58.7 | 81.303 |  |  |  | 1
| 1 | 2 | 66 | 65.9002 | 59.539 | 86.6 | 72 | 60 | 70 | 60 | 70 | 60 | 79.7776 | 87 | 70.6 | 101 | 57.4 | 113 | 59.2 | 82.34 |  |  |  | 24
| 1 | 3 | 65 | 66.4856 | 60.902 | 86.7 | 70 | 60 | 70 | 60 | 60 | 60 | 68.1764 | 87 | 66.2 | 105 | 56 | 105 | 59.2 | 78.304 |  |  |  | 16
| 1 | 4 | 65 | 65.084 | 56 | 89 | 71 | 70 | 70 | 60 | 70 | 60 | 69.0412 | 88 | 65.1 | 104 | 59 | 115 | 58.7 | 74.619 |  |  |  | 68.1944
| 1 | 5 | 65 | 65.3483 | 58.849 | 88.4 | 70 | 60 | 70 | 60 | 70 | 60 | 78.3356 | 87 | 69.1 | 102 | 55.1 | 107 | 58.7 | 72.756 |  |  |  | 78
| 1 | 6 | 65 | 66.6494 | 59.469 | 82.7 | 68 | 60 | 70 | 60 | 70 | 60 | 71.3179 | 85 | 70.9 | 98 | 60 | 132 | 60.2 | 79.577
| Mean |  | 65.1667 | 65.7743 | 58.9827 | 87.2333 | 70.1667 | 61.6667 | 70 | 60 | 68.3333 | 60 | 73.3571 | 86.8333 | 68.7833 | 101.8333 | 56.95 | 113.3333 | 59.1167 | 78.1498
| Median |  | 65 | 65.6242 | 59.303 | 87.55 | 70 | 60 | 70 | 60 | 70 | 60 | 72.4059 | 87 | 69.85 | 101.5 | 56.7 | 110.5 | 58.95 | 78.9405
| Std Dev. |  | 0.40825 | 0.6784 | 1.6228 | 2.5835 | 1.3292 | 4.0825 | 0 | 0 | 4.0825 | 0 | 4.8094 | 0.98319 | 2.5373 | 2.4833 | 2.2625 | 9.8928 | 0.58452 | 3.7715
| Rel.Std.Dev. |  | 0.6265% | 1.031% | 2.751% | 2.962% | 1.894% | 6.62% | 0% | 0% | 5.974% | 0% | 6.556% | 1.132% | 3.689% | 2.439% | 3.973% | 8.729% | 0.9888% | 4.826%
| PDM3 |  | -4.44% | -3.549% | -13.51% | 27.92% | 2.892% | -9.572% | 2.648% | -12.02% | 0.2037% | -12.02% | 7.571% | 27.33% | 0.8636% | 49.33% | -16.49% | 66.19% | -13.31% | 14.6%
| Z-Score (Absolute) |  | 0.27 | 0.22 | 0.75 | 1.47 | 0.13 | 0.54 | 0.11 | 0.67 | 0.02 | 0.67 | 0.38 | 1.44 | 0.02 | 2.62 | 0.91 | 3.53 | 0.74 | 0.76
| Analytical results for Be in OREAS 263b (Certified Value 0.68 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.67 | 0.63048 | 0.376 | 0.8 | 0.66 | 0.65 | 0.7 | 0.66 | 0.69 | 0.62 | 0.75753 | 0.86 | 0.69 | < 2 | 0.726 | [6 empty] | 1
| 1 | 2 | 0.63 | 0.64958 | 0.396 | 0.7 | 0.69 | 0.66 | 0.7 | 0.66 | 0.71 | 0.62 | 0.74022 | 0.83 | 0.67 | < 2 | 0.733 | [6 empty] | 25
| 1 | 3 | 0.67 | 0.66958 | 0.397 | 0.7 | 0.64 | 0.63 | 0.7 | 0.66 | 0.68 | 0.64 | 0.64044 | 0.83 | 0.66 | < 2 | 0.708 | [6 empty] | 16
| 1 | 4 | 0.6 | 0.622 |  | 0.7 | 0.7 | 0.68 | 0.69 | 0.64 | 0.7 | 0.6 | 0.66457 | 0.79 | 0.79 | < 2 | 0.702 | [6 empty] | 0.68066
| 1 | 5 | 0.8 | 0.67994 | 0.395 | 0.7 | 0.67 | 0.65 | 0.69 | 0.65 | 0.71 | 0.63 | 0.69936 | 0.84 | 0.69 | < 2 | 0.728 | [6 empty] | 79
| 1 | 6 | 0.8 | 0.69323 | 0.384 | 0.7 | 0.71 | 0.65 | 0.69 | 0.65 | 0.7 | 0.59 | 0.70458 | 0.84 | 0.82 | < 2 | 0.715
| Mean |  | 0.695 | 0.65747 | 0.3896 | 0.71667 | 0.67833 | 0.65333 | 0.695 | 0.65333 | 0.69833 | 0.61667 | 0.70112 | 0.83167 | 0.72 |  | 0.71867
| Median |  | 0.67 | 0.65958 | 0.395 | 0.7 | 0.68 | 0.65 | 0.695 | 0.655 | 0.7 | 0.62 | 0.70197 | 0.835 | 0.69 |  | 0.7205
| Std Dev. |  | 0.085499 | 0.028209 | 0.0092358 | 0.040825 | 0.026394 | 0.01633 | 0.0054772 | 0.008165 | 0.01169 | 0.018619 | 0.044146 | 0.023166 | 0.067528 |  | 0.012258
| Rel.Std.Dev. |  | 12.3% | 4.291% | 2.371% | 5.696% | 3.891% | 2.499% | 0.7881% | 1.25% | 1.674% | 3.019% | 6.297% | 2.785% | 9.379% |  | 1.706%
| PDM3 |  | 2.107% | -3.407% | -42.76% | 5.29% | -0.3417% | -4.015% | 2.107% | -4.015% | 2.597% | -9.402% | 3.005% | 22.19% | 5.78% |  | 5.584%
| Z-Score (Absolute) |  | 0 | 0.81 | 6.59 | NA | 0.36 | 0.9 | 0 | 0.9 | 0.07 | 1.69 | 0.13 | 2.95 | 0.54 | 6.58 | 0.51
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 263b (Certified Value 9.1 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 8.79 | 8.3728 | 11.7052 | 10.2 | 9.54 | 8.55 | 9 | 9.54 | 8.94 | 8.76 | 9.0135 | 8.97 | 9.77 | 9.11 | 10 | 8.96 | 11.3 | < 5 | 9.4 | 3.3 |  | 1
| 1 | 2 | 8.73 | 8.1809 | 11.9276 | 10.1 | 10.19 | 8.62 | 9.15 | 8.94 | 8.46 | 8.69 | 9.0355 | 8.97 | 9.68 | 8.68 | 10 | 9.01 | 10.3 | < 5 | 9.3 | 4.672 |  | 26
| 1 | 3 | 9.08 | 8.632 | 12.4699 | 9.64 | 9.61 | 8.05 | 9.04 | 8.48 | 8.31 | 8.86 | 9.497 | 8.97 | 9.79 | 8.52 | 10 | 8.97 | 11.1 | < 5 | 9.1 | 4.752 |  | 16
| 1 | 4 | 8.82 | 8.698 | 11.588 | 10.1 | 9.96 | 8.61 | 9.41 | 8.78 | 8.4 | 8.46 | 9.2652 | 9.97 | 9.45 | 9.71 | 10 | 9.38 | 10.3 | < 5 | 9.6 | 5.473 |  | 9.0964
| 1 | 5 | 9.17 | 8.357 | 11.3361 | 10 | 9.68 | 8.78 | 9.31 | 9.16 | 9.14 | 9.37 | 9.5686 | 7.97 | 9.7 | 8.44 | 10 | 8.89 | 11 | < 5 | 9.2 | 5.106 |  | 80
| 1 | 6 | 8.8 | 8.6416 | 12.7649 | 9.96 | 9.56 | 9.08 | 9.19 | 9.22 | 9.03 | 8.53 | 9.4069 | 8.97 | 9.66 | 8.57 | 10 | 9.03 | 10.8 | < 5 | 9.5 | 4.389
| Mean |  | 8.8983 | 8.4804 | 11.9653 | 10 | 9.7567 | 8.615 | 9.1833 | 9.02 | 8.7133 | 8.7783 | 9.2978 | 8.97 | 9.675 | 8.8383 | 10 | 9.04 | 10.8 |  | 9.35 | 4.6153
| Median |  | 8.81 | 8.5024 | 11.8164 | 10.05 | 9.645 | 8.615 | 9.17 | 9.05 | 8.7 | 8.725 | 9.3361 | 8.97 | 9.69 | 8.625 | 10 | 8.99 | 10.9 |  | 9.35 | 4.712
| Std Dev. |  | 0.18038 | 0.20632 | 0.54797 | 0.19555 | 0.26158 | 0.33673 | 0.15667 | 0.3703 | 0.36297 | 0.32493 | 0.23478 | 0.63246 | 0.12145 | 0.48799 | 0 | 0.17344 | 0.41952 |  | 0.18708 | 0.7456
| Rel.Std.Dev. |  | 2.027% | 2.433% | 4.58% | 1.956% | 2.681% | 3.909% | 1.706% | 4.105% | 4.166% | 3.701% | 2.525% | 7.051% | 1.255% | 5.521% | 0% | 1.919% | 3.884% |  | 2.001% | 16.15%
| PDM3 |  | -2.178% | -6.772% | 31.54% | 9.933% | 7.258% | -5.292% | 0.9556% | -0.84% | -4.211% | -3.497% | 2.214% | -1.39% | 6.361% | -2.837% | 9.933% | -0.6201% | 18.73% |  | 2.788% | -49.26%
| Z-Score (Absolute) |  | 0.25 | 1.1 | 6.02 | 2 | 1.51 | 0.83 | 0.33 | 0 | 0.63 | 0.49 | 0.57 | 0.1 | 1.34 | 0.37 | NA | 0.04 | 3.64 | 13.32 | 0.67 | 9
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 263b (Certified Value 0.162 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 0.15 | 0.15017 | 0.15912 | 0.18 | 0.17 | 0.16 | 0.17 | 0.16 | 0.16 | 0.16 | 0.16846 | 0.1637 | 0.17 | 0.16 | 0.19 | 0.14 | 0.16 | 0.16 | 0.16 | 0.14065 |  | 1
| 1 | 2 | 0.16 | 0.15328 | 0.16233 | 0.21 | 0.16 | 0.16 | 0.18 | 0.16 | 0.16 | 0.16 | 0.16095 | 0.1598 | 0.17 | 0.16 | 0.19 | 0.15 | 0.158 | 0.16 | 0.16 | 0.17228 |  | #N/A
| 1 | 3 | 0.16 | 0.15225 | 0.16197 | 0.18 | 0.16 | 0.16 | 0.17 | 0.17 | 0.16 | 0.17 | 0.16317 | 0.1661 | 0.17 | 0.16 | 0.19 | 0.15 | 0.156 | 0.16 | 0.16 | 0.16652 |  | 16
| 1 | 4 | 0.15 | 0.15473 | 0.1521 | 0.18 | 0.16 | 0.17 | 0.17 | 0.16 | 0.16 | 0.16 | 0.16442 | 0.1624 | 0.17 | 0.16 | 0.19 | 0.15 | 0.167 | 0.16 | 0.16 | 0.1399 |  | 0.16193
| 1 | 5 | 0.16 | 0.15609 | 0.16076 | 0.18 | 0.16 | 0.16 | 0.17 | 0.16 | 0.16 | 0.16 | 0.17014 | 0.1645 | 0.17 | 0.16 | 0.19 | 0.15 | 0.175 | 0.16 | 0.16 | 0.13928 |  | 81
| 1 | 6 | 0.16 | 0.15307 | 0.15849 | 0.16 | 0.17 | 0.16 | 0.17 | 0.16 | 0.16 | 0.16 | 0.16834 | 0.1644 | 0.17 | 0.16 | 0.18 | 0.15 | 0.173 | 0.17 | 0.16 | 0.15968
| Mean |  | 0.15667 | 0.15326 | 0.15913 | 0.18167 | 0.16333 | 0.16167 | 0.17167 | 0.16167 | 0.16 | 0.16167 | 0.16591 | 0.16348 | 0.17 | 0.16 | 0.18833 | 0.14833 | 0.16483 | 0.16167 | 0.16 | 0.15305
| Median |  | 0.16 | 0.15318 | 0.15994 | 0.18 | 0.16 | 0.16 | 0.17 | 0.16 | 0.16 | 0.16 | 0.16638 | 0.16405 | 0.17 | 0.16 | 0.19 | 0.15 | 0.1635 | 0.16 | 0.16 | 0.15017
| Std Dev. |  | 0.005164 | 0.002039 | 0.0037633 | 0.016021 | 0.005164 | 0.0040825 | 0.0040825 | 0.0040825 | 0 | 0.0040825 | 0.0035976 | 0.0021683 | 3.0405e-17 | 0 | 0.0040825 | 0.0040825 | 0.0080353 | 0.0040825 | 0 | 0.014911
| Rel.Std.Dev. |  | 3.296% | 1.33% | 2.365% | 8.819% | 3.162% | 2.525% | 2.378% | 2.525% | 0% | 2.525% | 2.168% | 1.326% | 1.789e-14% | 0% | 2.168% | 2.752% | 4.875% | 2.525% | 0% | 9.742%
| PDM3 |  | -3.25% | -5.351% | -1.729% | 12.19% | 0.8675% | -0.1618% | 6.014% | -0.1618% | -1.191% | -0.1618% | 2.461% | 0.9601% | 4.985% | -1.191% | 16.31% | -8.396% | 1.794% | -0.1618% | -1.191% | -5.481%
| Z-Score (Absolute) |  | 1.18 | 1.99 | 0.6 | 4.73 | 0.39 | 0 | 2.36 | 0 | 0.39 | 0 | 1 | 0.43 | 1.97 | 0.39 | 6.3 | 3.15 | 0.75 | 0 | 0.39 | 2.04
| Analytical results for Cd in OREAS 263b (Certified Value 0.035 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  |  | 3
| 1 | 1 | 0.04 | < 0.2 | 0.04205 | 0.05 | 0.07 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | < 0.05 | < 0.05 | 0.05 | 0.03 | 0.03 | 2.7 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 0.03 | < 0.2 | 0.044054 | 0.15 | 0.07 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | < 0.05 | < 0.05 | 0.05 | 0.03 | 0.03 | 2.3 | < 0.1 |  |  |  |  | 27
| 1 | 3 | 0.03 | < 0.2 | 0.038689 | 0.05 | 0.08 | 0.03 | 0.03 | 0.02 | 0.03 | 0.04 | < 0.05 | < 0.05 | 0.05 | 0.04 | 0.03 | 2.4 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 0.04 | < 0.2 | 0.035 | 0.07 | 0.08 | 0.04 | 0.03 | 0.03 | 0.04 | 0.05 | < 0.05 | < 0.05 | 0.03 | 0.03 | 0.03 | 2.6 | < 0.1 |  |  |  |  | 0.034871
| 1 | 5 | 0.03 | < 0.2 | 0.048069 | 0.04 | 0.07 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.03 | 0.03 | 2.5 | < 0.1 |  |  |  |  | 82
| 1 | 6 | 0.04 | < 0.2 | 0.044412 | 0.15 | 0.08 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | < 0.05 | < 0.05 | 0.05 | 0.03 | 0.03 | 2.5 | < 0.1
| Mean |  | 0.035 |  | 0.042046 | 0.085 | 0.075 | 0.031667 | 0.03 | 0.028333 | 0.036667 | 0.038333 |  |  | 0.045 | 0.031667 | 0.03 | 2.5
| Median |  | 0.035 |  | 0.043052 | 0.06 | 0.075 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 |  |  | 0.05 | 0.03 | 0.03 | 2.5
| Std Dev. |  | 0.0054772 |  | 0.0046222 | 0.051284 | 0.0054772 | 0.0040825 | 0 | 0.0040825 | 0.005164 | 0.0075277 |  |  | 0.0083666 | 0.0040825 | 0 | 0.14142
| Rel.Std.Dev. |  | 15.65% |  | 10.99% | 60.33% | 7.303% | 12.89% | 0% | 14.41% | 14.08% | 19.64% |  |  | 18.59% | 12.89% | 0% | 5.657%
| PDM3 |  | 0.3693% |  | 20.57% | 143.8% | 115.1% | -9.19% | -13.97% | -18.75% | 5.149% | 9.928% |  |  | 29.05% | -9.19% | -13.97% | 7069%
| Z-Score (Absolute) |  | 0.13 | 5.12 | 0.44 | 3.91 | 3.1 | 0.4 | 0.54 | 0.67 | 0 | 0.13 | 0.94 | 0.94 | 0.67 | 0.4 | 0.54 | 199.33 | 1.08
| Analytical results for Ce in OREAS 263b (Certified Value 48.2 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.21 | 1.28 | 1.30 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 01g | 0.1g | 40g | 0.5g | [5 empty] | 1
| 1 | 1 | 44.445 | 44.0817 | 68.3236 | 54.8 | 48.8 | 43.9 | 48.1 | 42.7 | 43.4 | 41.2 | 49.0135 | 55.8 | 52 | 76.2 | 60.4 | 47.833 | [5 empty] | 1
| 1 | 2 | 45.984 | 46.0755 | 68.4304 | 53.3 | 48.95 | 44.9 | 48.7 | 40.9 | 44 | 44 | 49.7625 | 54.2 | 53 | 77.6 | 58.8 | 47.572 | [5 empty] | 28
| 1 | 3 | 46.273 | 45.0045 | 68.432 | 53.9 | 48.08 | 42.9 | 49.3 | 42.9 | 42.3 | 41.8 | 47.2792 | 56 | 53 | 71.8 | 60.6 | 44.581 | [5 empty] | 16
| 1 | 4 | 45.215 | 44.4054 | 68.4176 | 55.5 | 49.41 | 45.8 | 48.4 | 40.7 | 44.1 | 41.4 | 47.5923 | 55.5 | 53 | 77 | 63 | 43.649 | [5 empty] | 48.1594
| 1 | 5 | 45.824 | 43.9717 | 68.3332 | 53.8 | 48.43 | 42.6 | 47.4 | 41 | 41.3 | 42.4 | 48.288 | 55.8 | 52 | 70.5 | 59.8 | 43.776 | [5 empty] | 83
| 1 | 6 | 46.392 | 45.6418 | 68.4764 | 52.1 | 48.46 | 41.7 | 48.5 | 41.7 | 44.7 | 40.8 | 49.838 | 56 | 51 | 78.6 | 58.2 | 45.762
| Mean |  | 45.6888 | 44.8634 | 68.4022 | 53.9 | 48.6883 | 43.6333 | 48.4 | 41.65 | 43.3 | 41.9333 | 48.6289 | 55.55 | 52.3333 | 75.2833 | 60.1333 | 45.5288
| Median |  | 45.904 | 44.7049 | 68.424 | 53.85 | 48.63 | 43.4 | 48.45 | 41.35 | 43.7 | 41.6 | 48.6507 | 55.8 | 52.5 | 76.6 | 60.1 | 45.1715
| Std Dev. |  | 0.73658 | 0.8615 | 0.060605 | 1.1815 | 0.46705 | 1.5306 | 0.63246 | 0.95446 | 1.2728 | 1.1501 | 1.0869 | 0.68629 | 0.8165 | 3.3217 | 1.6813 | 1.8462
| Rel.Std.Dev. |  | 1.612% | 1.92% | 0.0886% | 2.192% | 0.9593% | 3.508% | 1.307% | 2.292% | 2.939% | 2.743% | 2.235% | 1.235% | 1.56% | 4.412% | 2.796% | 4.055%
| PDM3 |  | -5.13% | -6.844% | 42.03% | 11.92% | 1.098% | -9.398% | 0.4996% | -13.52% | -10.09% | -12.93% | 0.9749% | 15.35% | 8.667% | 56.32% | 24.86% | -5.462%
| Z-Score (Absolute) |  | 0.38 | 0.49 | 2.66 | 0.72 | 0.02 | 0.65 | 0.02 | 0.92 | 0.7 | 0.88 | 0.02 | 0.94 | 0.51 | 3.58 | 1.55 | 0.4
| Analytical results for Co in OREAS 263b (Certified Value 3.36 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 3.4 | 3.2469 | 3.103 | 4.2 | 3.5 | 3 | 3.4 | 3.2 | 3.5 | 2.9 | 3.795 | 3.8 | 3.3 | 4 | 3.2 | 3.8 | 1.5 | 3 | 2.865 |  |  | 1
| 1 | 2 | 3.5 | 3.2557 | 3.005 | 4.3 | 3.5 | 3.2 | 3.4 | 3.1 | 3.6 | 3.2 | 4.2243 | 3.8 | 3.5 | 4 | 3.3 | 3.8 | 1.4 | 3 | 2.734 |  |  | 14
| 1 | 3 | 3.5 | 3.3562 | 3.121 | 4 | 3.6 | 3 | 3.4 | 3.3 | 3.4 | 2.9 | 3.9703 | 3.8 | 3.2 | 4 | 3.3 | 3.5 | 1.5 | 3 | 2.344 |  |  | 16
| 1 | 4 | 3.4 | 3.2953 | 3 | 4.1 | 3.3 | 3.3 | 3.4 | 3.2 | 3.5 | 3 | 3.7044 | 3.7 | 3.4 | 4 | 3.4 | 3.9 | 1.2 | 3.2 | 2.168 |  |  | 3.3635
| 1 | 5 | 3.5 | 3.3919 | 3.023 | 4.4 | 3.5 | 3.1 | 3.3 | 3.1 | 3.3 | 3.1 | 4.1766 | 3.7 | 3.2 | 4 | 3.3 | 3.75 | 1.5 | 3 | 2.332 |  |  | 84
| 1 | 6 | 3.4 | 3.4115 | 2.901 | 3.8 | 3.5 | 3.1 | 3.4 | 3.2 | 3.5 | 2.7 | 3.8962 | 3.8 | 3.2 | 4 | 3.4 | 3.99 | 1.5 | 3 | 2.422
| Mean |  | 3.45 | 3.3262 | 3.0255 | 4.1333 | 3.4833 | 3.1167 | 3.3833 | 3.1833 | 3.4667 | 2.9667 | 3.9611 | 3.7667 | 3.3 | 4 | 3.3167 | 3.79 | 1.4333 | 3.0333 | 2.4775
| Median |  | 3.45 | 3.3257 | 3.014 | 4.15 | 3.5 | 3.1 | 3.4 | 3.2 | 3.5 | 2.95 | 3.9332 | 3.8 | 3.25 | 4 | 3.3 | 3.8 | 1.5 | 3 | 2.383
| Std Dev. |  | 0.054772 | 0.07029 | 0.079629 | 0.21602 | 0.098319 | 0.1169 | 0.040825 | 0.075277 | 0.10328 | 0.17512 | 0.20664 | 0.05164 | 0.12649 | 0 | 0.075277 | 0.16613 | 0.12111 | 0.08165 | 0.26601
| Rel.Std.Dev. |  | 1.588% | 2.113% | 2.632% | 5.226% | 2.823% | 3.751% | 1.207% | 2.365% | 2.979% | 5.903% | 5.217% | 1.371% | 3.833% | 0% | 2.27% | 4.383% | 8.449% | 2.692% | 10.74%
| PDM3 |  | 2.57% | -1.109% | -10.05% | 22.89% | 3.561% | -7.34% | 0.5882% | -5.358% | 3.066% | -11.8% | 17.77% | 11.98% | -1.889% | 18.92% | -1.394% | 12.68% | -57.39% | -9.817% | -26.34%
| Z-Score (Absolute) |  | 0.35 | 0.01 | 0.81 | 2.22 | 0.44 | 0.56 | 0.17 | 0.38 | 0.4 | 0.97 | 1.75 | 1.22 | 0.06 | NA | 0.01 | 1.28 | 5.17 | 0.79 | 2.31
| Analytical results for Cr in OREAS 263b (Certified Value 33.9 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 40g | 0.5g |  |  | 2
| 1 | 1 | 33 | 32.6033 | 30.928 | 38.37 | 35 | 33 | 33 | 35 | 34 | 34 | 34 | 35.8091 | 36 | 34.4 | 39 | 31.6 | 33 | 30 | 32.254 |  |  | 1
| 1 | 2 | 33 | 31.8154 | 31.733 | 34.58 | 35 | 34 | 34 | 35 | 34 | 34 | 34 | 34.9021 | 36 | 34.5 | 39 | 33.1 | 31 | 30 | 32.935 |  |  | 30
| 1 | 3 | 34 | 33.7127 | 32.228 | 38.87 | 37 | 34 | 33 | 35 | 34 | 34 | 35 | 34.5231 | 36 | 34.2 | 39 | 32.8 | 29 | 30 | 33.015 |  |  | 16
| 1 | 4 | 33 | 33.1163 | 30 | 37.39 | 35 | 34 | 35 | 36 | 33 | 34 | 32 | 33.5473 | 37 | 33.7 | 41 | 33.9 | 36 | 30 | 31.886 |  |  | 33.9228
| 1 | 5 | 33 | 33.3533 | 31.478 | 38.03 | 35 | 34 | 33 | 35 | 34 | 34 | 35 | 33.9285 | 37 | 33.2 | 38 | 32.8 | 38 | 30 | 32.747 |  |  | 85
| 1 | 6 | 34 | 33.0762 | 31.343 | 38.5 | 33 | 34 | 34 | 35 | 34 | 35 | 33 | 35.9874 | 37 | 34.4 | 38 | 34 | 39 | 30 | 34.071
| Mean |  | 33.3333 | 32.9462 | 31.285 | 37.6233 | 35 | 33.8333 | 33.6667 | 35.1667 | 33.8333 | 34.1667 | 33.8333 | 34.7829 | 36.5 | 34.0667 | 39 | 33.0333 | 34.3333 | 30 | 32.818
| Median |  | 33 | 33.0962 | 31.4105 | 38.2 | 35 | 34 | 33.5 | 35 | 34 | 34 | 34 | 34.7126 | 36.5 | 34.3 | 39 | 32.95 | 34.5 | 30 | 32.841
| Std Dev. |  | 0.5164 | 0.66262 | 0.76279 | 1.5724 | 1.2649 | 0.40825 | 0.8165 | 0.40825 | 0.40825 | 0.40825 | 1.169 | 0.98409 | 0.54772 | 0.51251 | 1.0954 | 0.87788 | 3.9833 | 0 | 0.75024
| Rel.Std.Dev. |  | 1.549% | 2.011% | 2.438% | 4.179% | 3.614% | 1.207% | 2.425% | 1.161% | 1.207% | 1.195% | 3.455% | 2.829% | 1.501% | 1.504% | 2.809% | 2.658% | 11.6% | 0% | 2.286%
| PDM3 |  | -1.738% | -2.879% | -7.776% | 10.91% | 3.175% | -0.2639% | -0.7552% | 3.667% | -0.2639% | 0.7188% | -0.2639% | 2.535% | 7.597% | 0.424% | 14.97% | -2.622% | 1.21% | -11.56% | -3.257%
| Z-Score (Absolute) |  | 0.38 | 0.67 | 1.94 | 2.88 | 0.89 | 0 | 0.13 | 1.01 | 0 | 0.25 | 0 | 0.72 | 2.03 | 0.18 | 3.93 | 0.61 | 0.38 | 2.91 | 0.77
| Analytical results for Cs in OREAS 263b (Certified Value 2.09 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 2
| 1 | 1 | 2.5 | 2.3335 | 2.37 | 1.96 | 1.98 | 2.11 | 1.86 | 1.81 | 2.01 | 2.3232 | 3.19 | 3.7 | 1.8 | [8 empty] | 1
| 1 | 2 | 2.54 | 2.3726 | 2.17 | 1.96 | 2.03 | 2.1 | 1.81 | 1.93 | 2.15 | 2.2944 | 3.25 | 3.8 | 1.8 | [8 empty] | 31
| 1 | 3 | 2.55 | 2.3293 | 2.28 | 1.94 | 1.89 | 2.08 | 1.88 | 1.81 | 2.06 | 2.1949 | 3.3 | 3.7 | 1.8 | [8 empty] | 16
| 1 | 4 | 2.48 | 2.3461 | 2.35 | 2.02 | 2.02 | 2.12 | 1.79 | 1.91 | 2.04 | 2.1874 | 3.24 | 4 | 1.9 | [8 empty] | 2.0892
| 1 | 5 | 2.53 | 2.35 | 2.21 | 2.05 | 1.93 | 2.09 | 1.79 | 1.86 | 2 | 2.2289 | 3.27 | 3.6 | 1.8 | [8 empty] | 86
| 1 | 6 | 2.53 | 2.378 | 2.17 | 1.98 | 1.91 | 2.12 | 1.83 | 1.94 | 2.02 | 2.3055 | 3.32 | 4.3 | 1.8
| Mean |  | 2.5217 | 2.3516 | 2.2583 | 1.985 | 1.96 | 2.1033 | 1.8267 | 1.8767 | 2.0467 | 2.2557 | 3.2617 | 3.85 | 1.8167
| Median |  | 2.53 | 2.348 | 2.245 | 1.97 | 1.955 | 2.105 | 1.82 | 1.885 | 2.03 | 2.2617 | 3.26 | 3.75 | 1.8
| Std Dev. |  | 0.026394 | 0.019994 | 0.088638 | 0.041833 | 0.058652 | 0.01633 | 0.037238 | 0.058538 | 0.055015 | 0.059359 | 0.046224 | 0.25884 | 0.040825
| Rel.Std.Dev. |  | 1.047% | 0.8502% | 3.925% | 2.107% | 2.992% | 0.7764% | 2.039% | 3.119% | 2.688% | 2.631% | 1.417% | 6.723% | 2.247%
| PDM3 |  | 20.7% | 12.56% | 8.094% | -4.989% | -6.186% | 0.6746% | -12.57% | -10.17% | -2.038% | 7.968% | 56.12% | 84.28% | -13.05%
| Z-Score (Absolute) |  | 1.24 | 0.74 | 0.46 | 0.35 | 0.43 | 0 | 0.82 | 0.67 | 0.17 | 0.45 | 3.45 | 5.2 | 0.85
| Analytical results for Cu in OREAS 263b (Certified Value 73 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 1
| 1 | 1 | 72.1 | 71.2811 | 75.69 | 68.15 | 77 | 71 | 71.2 | 73.2 | 74.2 | 74.1 | 69.7 | 72.5293 | 79 | 70 | 73.72 | 71 | 68.17 | 68.3 | 78.7 | 61 | 74.622 | 1
| 1 | 2 | 72.6 | 71.1976 | 75.812 | 63.66 | 79.2 | 74 | 74.6 | 71.8 | 72.3 | 77 | 72 | 72.3732 | 76 | 68 | 75.43 | 73 | 69.33 | 65.6 | 79.3 | 61 | 74.307 | 16
| 1 | 3 | 72.2 | 73.9813 | 75.23 | 69.24 | 71.5 | 74 | 71.5 | 72.9 | 72.8 | 75.4 | 73.8 | 72.9355 | 79 | 71 | 71.97 | 74 | 70.16 | 65.2 | 78.8 | 62 | 75.764 | 16
| 1 | 4 | 71.6 | 73.6076 | 76 | 67.33 | 72.7 | 74 | 74 | 72.9 | 72.5 | 78.3 | 71.6 | 71.9185 | 76 | 71 | 72.64 | 73 | 69.71 | 67.4 | 78.1 | 62 | 74.545 | 72.7162
| 1 | 5 | 72.5 | 76.7786 | 77.134 | 69 | 74.6 | 76 | 70.7 | 71.9 | 71 | 75.1 | 74.9 | 73.3476 | 79 | 71 | 70.32 | 72 | 70.31 | 85.8 | 78.5 | 63 | 72.964 | 87
| 1 | 6 | 72.8 | 74.7805 | 73.536 | 71.5 | 71.9 | 71 | 72.1 | 71.8 | 73.9 | 77.6 | 68.1 | 73.6877 | 80 | 71 | 70.56 | 71 | 71.22 | 67.5 | 78.4 | 62 | 75.626
| Mean |  | 72.3 | 73.6044 | 75.567 | 68.1467 | 74.4833 | 73.3333 | 72.35 | 72.4167 | 72.7833 | 76.25 | 71.6833 | 72.7986 | 78.1667 | 70.3333 | 72.44 | 72.3333 | 69.8167 | 69.9667 | 78.6333 | 61.8333 | 74.638
| Median |  | 72.35 | 73.7945 | 75.751 | 68.575 | 73.65 | 74 | 71.8 | 72.4 | 72.65 | 76.2 | 71.8 | 72.7324 | 79 | 71 | 72.305 | 72.5 | 69.935 | 67.45 | 78.6 | 62 | 74.5835
| Std Dev. |  | 0.42895 | 2.135 | 1.1795 | 2.6064 | 3.0773 | 1.9664 | 1.5884 | 0.64936 | 1.1618 | 1.6282 | 2.5183 | 0.65404 | 1.7224 | 1.2111 | 1.9426 | 1.2111 | 1.0286 | 7.8477 | 0.40825 | 0.75277 | 1.0163
| Rel.Std.Dev. |  | 0.5933% | 2.901% | 1.561% | 3.825% | 4.132% | 2.681% | 2.195% | 0.8967% | 1.596% | 2.135% | 3.513% | 0.8984% | 2.203% | 1.722% | 2.682% | 1.674% | 1.473% | 11.22% | 0.5192% | 1.217% | 1.362%
| PDM3 |  | -0.5724% | 1.221% | 3.92% | -6.284% | 2.43% | 0.8486% | -0.5037% | -0.412% | 0.09225% | 4.86% | -1.42% | 0.1133% | 7.495% | -3.277% | -0.3799% | -0.5266% | -3.988% | -3.781% | 8.137% | -14.97% | 2.643%
| Z-Score (Absolute) |  | 0.05 | 0.38 | 1.03 | 1.42 | 0.67 | 0.29 | 0.03 | 0.01 | 0.11 | 1.26 | 0.25 | 0.12 | 1.89 | 0.7 | 0 | 0.04 | 0.87 | 0.82 | 2.04 | 3.5 | 0.73
| Analytical results for Dy in OREAS 263b (Indicative Value 1.66 ppm) | [22 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 2
| 1 | 1 | 1.594 | 1.6238 | 2.8396 | 1.9 | 1.54 | 0.722 | [15 empty] | 1
| 1 | 2 | 1.609 | 1.5942 | 2.7854 | 1.9 | 1.5 | 0.844 | [15 empty] | 4
| 1 | 3 | 1.629 | 1.6394 | 2.7974 | 1.9 | 1.53 | 0.838 | [15 empty] | 16
| 1 | 4 | 1.608 | 1.6177 | 2.8286 | 1.9 | 1.56 | 0.94 | [15 empty] | 1.6615
| 1 | 5 | 1.65 | 1.5952 | 2.8473 | 1.8 | 1.54 | 0.572 | [15 empty] | 10
| 1 | 6 | 1.689 | 1.6155 | 2.7894 | 1.8 | 1.54 | 0.863
| Mean |  | 1.6298 | 1.6143 | 2.8146 | 1.8667 | 1.535 | 0.7965
| Median |  | 1.619 | 1.6166 | 2.813 | 1.9 | 1.54 | 0.841
| Std Dev. |  | 0.034937 | 0.017335 | 0.027106 | 0.05164 | 0.019748 | 0.13034
| Rel.Std.Dev. |  | 2.144% | 1.074% | 0.963% | 2.766% | 1.287% | 16.36%
| PDM3 |  | -1.903% | -2.838% | 69.41% | 12.35% | -7.611% | -52.06%
| Z-Score (Absolute) |  | 0.03 | 0.03 | 4.85 | 0.99 | 0.35 | 3.36
| Analytical results for Er in OREAS 263b (Indicative Value 0.61 ppm) | [22 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.588 | 0.60336 | 1.1236 | 0.7 | 0.61 | [16 empty] | 1
| 1 | 2 | 0.604 | 0.61574 | 1.0654 | 0.7 | 0.6 | [16 empty] | 5
| 1 | 3 | 0.6 | 0.6152 | 1.0595 | 0.7 | 0.61 | [16 empty] | 16
| 1 | 4 | 0.608 | 0.61991 | 1.1293 | 0.7 | 0.63 | [16 empty] | 0.60794
| 1 | 5 | 0.613 | 0.60287 | 1.0796 | 0.7 | 0.61 | [16 empty] | 11
| 1 | 6 | 0.612 | 0.6009 | 1.0656 | 0.7 | 0.6
| Mean |  | 0.60417 | 0.60966 | 1.0872 | 0.7 | 0.61
| Median |  | 0.606 | 0.60928 | 1.0726 | 0.7 | 0.61
| Std Dev. |  | 0.0093041 | 0.0081889 | 0.031194 | 1.2162e-16 | 0.010954
| Rel.Std.Dev. |  | 1.54% | 1.343% | 2.869% | 1.737e-14% | 1.796%
| PDM3 |  | -0.6212% | 0.2829% | 78.83% | 15.14% | 0.3383%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 55.16 | 10.4 | 0
| Analytical results for Eu in OREAS 263b (Certified Value 0.6 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 3
| 1 | 1 | 0.582 | 0.56617 | 0.94456 | 0.7 | 0.552 | 0.551 | [15 empty] | 1
| 1 | 2 | 0.598 | 0.57403 | 0.94024 | 0.7 | 0.544 | 0.585 | [15 empty] | 35
| 1 | 3 | 0.608 | 0.58263 | 0.94672 | 0.7 | 0.548 | 0.562 | [15 empty] | 16
| 1 | 4 | 0.605 | 0.5726 | 1.0146 | 0.7 | 0.544 | 0.492 | [15 empty] | 0.59542
| 1 | 5 | 0.596 | 0.55177 | 0.97708 | 0.7 | 0.55 | 0.586 | [15 empty] | 88
| 1 | 6 | 0.621 | 0.59253 | 0.9596 | 0.7 | 0.554 | 0.545
| Mean |  | 0.60167 | 0.57329 | 0.96381 | 0.7 | 0.54867 | 0.5535
| Median |  | 0.6015 | 0.57332 | 0.95316 | 0.7 | 0.549 | 0.5565
| Std Dev. |  | 0.013095 | 0.013945 | 0.028269 | 1.2162e-16 | 0.0041312 | 0.034588
| Rel.Std.Dev. |  | 2.176% | 2.432% | 2.933% | 1.737e-14% | 0.7529% | 6.249%
| PDM3 |  | 1.048% | -3.717% | 61.87% | 17.56% | -7.853% | -7.041%
| Z-Score (Absolute) |  | 0.26 | 0.26 | 6.97 | 2.08 | 0.72 | 0.63
| Analytical results for Fe in OREAS 263b (Certified Value 2.29 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 2.27 | 2.2417 | 2.3198 | 2.45 | 2.29 | 2.17 | 2.23 | 2.24 | 2.28 | 2.15 | 2.3663 | 2.0779 | 2.42 | 2.32 | 2.35 | 2.07 | 2.33 | 2.27 | 2.42 | 2.3685 |  | 1
| 1 | 2 | 2.28 | 2.1841 | 2.3547 | 2.45 | 2.29 | 2.17 | 2.22 | 2.22 | 2.32 | 2.13 | 2.4079 | 2.0362 | 2.33 | 2.36 | 2.4 | 2.22 | 2.33 | 2.23 | 2.4 | 2.3731 |  | #N/A
| 1 | 3 | 2.33 | 2.2852 | 2.3543 | 2.33 | 2.31 | 2.13 | 2.23 | 2.28 | 2.28 | 2.19 | 2.328 | 2.1254 | 2.45 | 2.34 | 2.38 | 2.23 | 2.29 | 2.24 | 2.4 | 2.345 |  | 16
| 1 | 4 | 2.27 | 2.2703 | 2.2046 | 2.43 | 2.28 | 2.22 | 2.21 | 2.22 | 2.29 | 2.06 | 2.3072 | 2.082 | 2.39 | 2.28 | 2.39 | 2.24 | 2.32 | 2.22 | 2.41 | 2.3421 |  | 2.2855
| 1 | 5 | 2.28 | 2.3144 | 2.3308 | 2.44 | 2.34 | 2.16 | 2.21 | 2.24 | 2.32 | 2.15 | 2.3351 | 2.0933 | 2.37 | 2.27 | 2.36 | 2.19 | 2.4 | 2.19 | 2.44 | 2.3427 |  | 89
| 1 | 6 | 2.31 | 2.2945 | 2.3097 | 2.32 | 2.31 | 2.18 | 2.2 | 2.24 | 2.28 | 2.06 | 2.3748 | 2.1157 | 2.4 | 2.28 | 2.32 | 2.23 | 2.3 | 2.24 | 2.45 | 2.3667
| Mean |  | 2.29 | 2.265 | 2.3123 | 2.4033 | 2.3033 | 2.1717 | 2.2167 | 2.24 | 2.295 | 2.1233 | 2.3532 | 2.0884 | 2.3933 | 2.3083 | 2.3667 | 2.1967 | 2.3283 | 2.2317 | 2.42 | 2.3563
| Median |  | 2.28 | 2.2778 | 2.3253 | 2.435 | 2.3 | 2.17 | 2.215 | 2.24 | 2.285 | 2.14 | 2.3507 | 2.0877 | 2.395 | 2.3 | 2.37 | 2.225 | 2.325 | 2.235 | 2.415 | 2.3558
| Std Dev. |  | 0.024495 | 0.046561 | 0.055793 | 0.06121 | 0.021602 | 0.029269 | 0.012111 | 0.021909 | 0.019748 | 0.052789 | 0.036623 | 0.031668 | 0.041312 | 0.037103 | 0.029439 | 0.064395 | 0.038687 | 0.026394 | 0.020976 | 0.014497
| Rel.Std.Dev. |  | 1.07% | 2.056% | 2.413% | 2.547% | 0.9379% | 1.348% | 0.5463% | 0.9781% | 0.8605% | 2.486% | 1.556% | 1.516% | 1.726% | 1.607% | 1.244% | 2.931% | 1.662% | 1.183% | 0.8668% | 0.6152%
| PDM3 |  | 0.1958% | -0.8966% | 1.172% | 5.155% | 0.7792% | -4.982% | -3.013% | -1.992% | 0.4146% | -7.096% | 2.961% | -8.624% | 4.717% | 0.9979% | 3.55% | -3.888% | 1.873% | -2.357% | 5.884% | 3.098%
| Z-Score (Absolute) |  | 0.1 | 0.36 | 0.14 | 1.11 | 0.04 | 1.36 | 0.88 | 0.63 | 0.04 | 1.87 | 0.58 | 2.24 | 1 | 0.1 | 0.72 | 1.09 | 0.31 | 0.72 | 1.29 | 0.61
| Analytical results for Ga in OREAS 263b (Certified Value 2.89 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | [6 empty] | 3
| 1 | 1 | 2.84 | 3.862 | 3.31 | 3.08 | 2.82 | 2.86 | 2.67 | 2.64 | 2.73 | 3.1133 | 3.9 | 3.2 | 2.4 | 4.79 | < 5 | [6 empty] | 1
| 1 | 2 | 2.89 | 3.7495 | 3.62 | 2.99 | 2.89 | 2.85 | 2.54 | 2.79 | 2.87 | 3.1941 | 4 | 3.1 | 2.5 | 4.93 | < 5 | [6 empty] | 36
| 1 | 3 | 2.85 | 3.9248 | 3.42 | 2.9 | 2.72 | 2.89 | 2.69 | 2.76 | 2.79 | 2.9935 | 4.1 | 3 | 2.5 | 4.44 | < 5 | [6 empty] | 16
| 1 | 4 | 2.8 | 3.8363 | 3.47 | 2.97 | 2.96 | 2.87 | 2.57 | 2.89 | 2.69 | 2.9933 | 4.1 | 3.2 | 2.6 | 5.68 | < 5 | [6 empty] | 2.8928
| 1 | 5 | 2.85 | 3.8518 | 3.5 | 3.08 | 2.76 | 2.87 | 2.59 | 2.76 | 2.77 | 3.0435 | 4 | 3.1 | 2.5 | 5.13 | < 5 | [6 empty] | 90
| 1 | 6 | 2.86 | 3.9727 | 3.13 | 3.13 | 2.72 | 2.86 | 2.67 | 2.79 | 2.66 | 3.1539 | 4 | 3.1 | 2.6 | 5.93 | < 5
| Mean |  | 2.8483 | 3.8662 | 3.4083 | 3.025 | 2.8117 | 2.8667 | 2.6217 | 2.7717 | 2.7517 | 3.0819 | 4.0167 | 3.1167 | 2.5167 | 5.15
| Median |  | 2.85 | 3.8569 | 3.445 | 3.035 | 2.79 | 2.865 | 2.63 | 2.775 | 2.75 | 3.0784 | 4 | 3.1 | 2.5 | 5.03
| Std Dev. |  | 0.029269 | 0.076832 | 0.16987 | 0.085965 | 0.097656 | 0.013663 | 0.062743 | 0.080353 | 0.075476 | 0.084728 | 0.075277 | 0.075277 | 0.075277 | 0.56068
| Rel.Std.Dev. |  | 1.028% | 1.987% | 4.984% | 2.842% | 3.473% | 0.4766% | 2.393% | 2.899% | 2.743% | 2.749% | 1.874% | 2.415% | 2.991% | 10.89%
| PDM3 |  | -1.536% | 33.65% | 17.82% | 4.572% | -2.803% | -0.9017% | -9.371% | -4.186% | -4.877% | 6.54% | 38.85% | 7.741% | -13% | 78.03%
| Z-Score (Absolute) |  | 0.05 | 2.75 | 1.49 | 0.44 | 0.15 | 0 | 0.67 | 0.26 | 0.32 | 0.59 | 3.17 | 0.69 | 0.96 | 6.28 | 1.01
| Analytical results for Gd in OREAS 263b (Indicative Value 2.83 ppm) | [22 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 2.614 | 2.7765 | 3.2 | 2.57 | [17 empty] | 1
| 1 | 2 | 2.705 | 2.7451 | 3.3 | 2.74 | [17 empty] | 6
| 1 | 3 | 2.666 | 2.8365 | 3.2 | 2.52 | [17 empty] | 16
| 1 | 4 | 2.658 | 2.7804 | 3.1 | 2.67 | [17 empty] | 2.831
| 1 | 5 | 2.68 | 2.762 | 3.3 | 2.7 | [17 empty] | 12
| 1 | 6 | 2.774 | 2.8271 | 3.1 | 2.72
| Mean |  | 2.6828 | 2.7879 | 3.2 | 2.6533
| Median |  | 2.673 | 2.7785 | 3.2 | 2.685
| Std Dev. |  | 0.053757 | 0.036288 | 0.089443 | 0.088468
| Rel.Std.Dev. |  | 2.004% | 1.302% | 2.795% | 3.334%
| PDM3 |  | -5.235% | -1.522% | 13.03% | -6.277%
| Z-Score (Absolute) |  | 0.53 | 0.53 | 4.65 | 0.82
| Analytical results for Ge in OREAS 263b (Indicative Value 0.083 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.21 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.1g | [12 empty] | 3
| 1 | 1 | < 0.2 | < 0.1 | 0.23 | 0.05 | 0.06 | 0.08 | 0.06 | 0.11 | 0.12 | [12 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.24 | 0.05 | 0.05 | 0.08 | 0.06 | 0.12 | 0.11 | [12 empty] | 7
| 1 | 3 | < 0.2 | < 0.1 | 0.22 | 0.05 | 0.07 | 0.09 | 0.07 | 0.11 | 0.12 | [12 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.23 | 0.05 | 0.06 | 0.08 | 0.06 | 0.12 | 0.13 | [12 empty] | 0.0825
| 1 | 5 | < 0.2 | < 0.1 | 0.21 | 0.05 | 0.06 | 0.08 | 0.06 | 0.11 | 0.13 | [12 empty] | 13
| 1 | 6 | < 0.2 | < 0.1 | 0.22 | 0.05 | 0.06 | 0.08 | 0.07 | 0.12 | 0.14
| Mean |  |  |  | 0.225 | 0.05 | 0.06 | 0.081667 | 0.063333 | 0.115 | 0.125
| Median |  |  |  | 0.225 | 0.05 | 0.06 | 0.08 | 0.06 | 0.115 | 0.125
| Std Dev. |  |  |  | 0.010488 | 7.6012e-18 | 0.0063246 | 0.0040825 | 0.005164 | 0.0054772 | 0.010488
| Rel.Std.Dev. |  |  |  | 4.661% | 1.52e-14% | 10.54% | 4.999% | 8.154% | 4.763% | 8.39%
| PDM3 |  |  |  | 172.7% | -39.39% | -27.27% | -1.01% | -23.23% | 39.39% | 51.52%
| Z-Score (Absolute) |  | 0.39 | 0.67 | 3.05 | 0.67 | 0.46 | 0 | 0.39 | 0.71 | 0.92
| Analytical results for Hf in OREAS 263b (Certified Value 0.37 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 3
| 1 | 1 | 0.36 | 0.34277 | < 0.1 | 0.36 | 0.39 | 0.38 | 0.37 | 0.39 | 0.39 | 0.42277 | 0.29 | 0.4 | [9 empty] | 1
| 1 | 2 | 0.35 | 0.34538 | < 0.1 | 0.34 | 0.37 | 0.38 | 0.36 | 0.35 | 0.39 | 0.41225 | 0.32 | 0.37 | [9 empty] | 22
| 1 | 3 | 0.36 | 0.35512 | < 0.1 | 0.38 | 0.36 | 0.37 | 0.36 | 0.39 | 0.41 | 0.41548 | 0.28 | 0.34 | [9 empty] | 16
| 1 | 4 | 0.37 | 0.37107 | < 0.1 | 0.34 | 0.37 | 0.37 | 0.33 | 0.37 | 0.42 | 0.42521 | 0.33 | 0.42 | [9 empty] | 0.37185
| 1 | 5 | 0.38 | 0.36227 | < 0.1 | 0.37 | 0.37 | 0.37 | 0.34 | 0.37 | 0.38 | 0.41225 | 0.33 | 0.45 | [9 empty] | 91
| 1 | 6 | 0.39 | 0.3558 | 0.1 | 0.37 | 0.39 | 0.38 | 0.36 | 0.38 | 0.4 | 0.42557 | 0.34 | 0.48
| Mean |  | 0.36833 | 0.3554 | 0.1 | 0.36 | 0.375 | 0.375 | 0.35333 | 0.375 | 0.39833 | 0.41892 | 0.315 | 0.41
| Median |  | 0.365 | 0.35546 | 0.1 | 0.365 | 0.37 | 0.375 | 0.36 | 0.375 | 0.395 | 0.41913 | 0.325 | 0.41
| Std Dev. |  | 0.01472 | 0.010517 |  | 0.016733 | 0.012247 | 0.0054772 | 0.015055 | 0.015166 | 0.01472 | 0.006314 | 0.02429 | 0.051381
| Rel.Std.Dev. |  | 3.996% | 2.959% |  | 4.648% | 3.266% | 1.461% | 4.261% | 4.044% | 3.695% | 1.507% | 7.711% | 12.53%
| PDM3 |  | -0.9451% | -4.423% | -73.11% | -3.186% | 0.8478% | 0.8478% | -4.979% | 0.8478% | 7.123% | 12.66% | -15.29% | 10.26%
| Z-Score (Absolute) |  | 0.23 | 0.67 | NA | 0.52 | 0 | 0 | 0.75 | 0 | 0.8 | 1.51 | 2.06 | 1.2
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 263b (Certified Value 0.1 ppm) | [22 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.1g | 01g | 40g | [6 empty] | 2
| 1 | 1 | < 1 | 0.13 | 0.11 | 0.11 | 0.11 | 0.09 | 0.08 | 0.09 | 0.0519 | 0.18 | 0.096 | 0.092 | 0.13 | < 5 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 1 | 0.12 | 0.1 | 0.09 | 0.1 | 0.1 | 0.1 | 0.1 | 0.066224 | 0.17 | 0.08 | 0.095 | 0.15 | < 5 | < 0.5 | [6 empty] | 2
| 1 | 3 | < 1 | 0.16 | 0.12 | 0.09 | 0.09 | 0.1 | 0.1 | 0.12 | 0.049603 | 0.18 | 0.063 | 0.083 | 0.12 | < 5 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 1 | 0.12 | 0.13 | 0.1 | 0.1 | 0.09 | 0.09 | 0.11 | 0.070049 | 0.18 | 0.075 | 0.09 | 0.15 | < 5 | < 0.5 | [6 empty] | 0.10045
| 1 | 5 | < 1 | 0.14 | 0.11 | 0.09 | 0.08 | 0.1 | 0.12 | 0.11 | 0.064179 | 0.17 | 0.061 | 0.085 | 0.13 | < 5 | < 0.5 | [6 empty] | 92
| 1 | 6 | < 1 | 0.16 | 0.11 | 0.09 | 0.08 | 0.08 | 0.09 | 0.12 | 0.051723 | 0.19 | 0.088 | 0.098 | 0.16 | < 5 | < 0.5
| Mean |  |  | 0.13833 | 0.11333 | 0.095 | 0.093333 | 0.093333 | 0.096667 | 0.10833 | 0.058946 | 0.17833 | 0.077167 | 0.0905 | 0.14
| Median |  |  | 0.135 | 0.11 | 0.09 | 0.095 | 0.095 | 0.095 | 0.11 | 0.058039 | 0.18 | 0.0775 | 0.091 | 0.14
| Std Dev. |  |  | 0.018348 | 0.010328 | 0.0083666 | 0.012111 | 0.008165 | 0.013663 | 0.01169 | 0.0088627 | 0.0075277 | 0.013761 | 0.0057533 | 0.015492
| Rel.Std.Dev. |  |  | 13.26% | 9.113% | 8.807% | 12.98% | 8.748% | 14.13% | 10.79% | 15.04% | 4.221% | 17.83% | 6.357% | 11.07%
| PDM3 |  |  | 37.71% | 12.83% | -5.425% | -7.084% | -7.084% | -3.766% | 7.848% | -41.32% | 77.54% | -23.18% | -9.905% | 39.37%
| Z-Score (Absolute) |  | 8.8 | 0.67 | 0.11 | 0.3 | 0.34 | 0.34 | 0.26 | 0 | 1.11 | 1.57 | 0.7 | 0.4 | 0.71 | 53.76 | 3.18
| Analytical results for Ho in OREAS 263b (Certified Value 0.27 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 3
| 1 | 1 | 0.255 | 0.2464 | 0.29764 | 0.3 | 0.26 | [16 empty] | 1
| 1 | 2 | 0.249 | 0.2478 | 0.28516 | 0.3 | 0.255 | [16 empty] | 23
| 1 | 3 | 0.265 | 0.25231 | 0.2788 | 0.3 | 0.26 | [16 empty] | 16
| 1 | 4 | 0.26 | 0.24664 | 0.3142 | 0.3 | 0.255 | [16 empty] | 0.2708
| 1 | 5 | 0.258 | 0.25141 | 0.28684 | 0.3 | 0.255 | [16 empty] | 93
| 1 | 6 | 0.268 | 0.249 | 0.27268 | 0.3 | 0.255
| Mean |  | 0.25917 | 0.24892 | 0.28922 | 0.3 | 0.25667
| Median |  | 0.259 | 0.2484 | 0.286 | 0.3 | 0.255
| Std Dev. |  | 0.0068532 | 0.0024707 | 0.014827 | 0 | 0.002582
| Rel.Std.Dev. |  | 2.644% | 0.9925% | 5.127% | 0% | 1.006%
| PDM3 |  | -4.294% | -8.077% | 6.804% | 10.78% | -5.218%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.98 | 2.69 | 0.16
| Analytical results for In in OREAS 263b (Indicative Value 0.024 ppm) | [22 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.03 | 0.094 | 0.025 | 0.023 | 0.024 | 0.024 | 0.025 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.03 | 0.093 | 0.025 | 0.024 | 0.022 | 0.025 | 0.025 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 4
| 1 | 3 | 0.02 | 0.03 | 0.098 | 0.024 | 0.024 | 0.024 | 0.022 | 0.025 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 16
| 1 | 4 | 0.02 | 0.03 | 0.093 | 0.025 | 0.021 | 0.024 | 0.023 | 0.026 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 0.0235
| 1 | 5 | 0.03 | 0.03 | 0.091 | 0.023 | 0.023 | 0.022 | 0.022 | 0.023 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 10
| 1 | 6 | 0.02 | 0.03 | 0.098 | 0.023 | 0.023 | 0.021 | 0.022 | 0.024 | < 0.05 | < 0.05 | 0.03 | < 0.5
| Mean |  | 0.023333 | 0.03 | 0.0945 | 0.024167 | 0.023 | 0.022833 | 0.023 | 0.024667 |  |  | 0.03
| Median |  | 0.02 | 0.03 | 0.0935 | 0.0245 | 0.023 | 0.023 | 0.0225 | 0.025 |  |  | 0.03
| Std Dev. |  | 0.005164 | 0 | 0.002881 | 0.00098319 | 0.0010954 | 0.0013292 | 0.0012649 | 0.0010328 |  |  | 0
| Rel.Std.Dev. |  | 22.13% | 0% | 3.049% | 4.068% | 4.763% | 5.821% | 5.5% | 4.187% |  |  | 0%
| PDM3 |  | -0.7092% | 27.66% | 302.1% | 2.837% | -2.128% | -2.837% | -2.128% | 4.965% |  |  | 27.66%
| Z-Score (Absolute) |  | 0.55 | 1.9 | 25.62 | 0.25 | 0.67 | 0.74 | 0.67 | 0.06 | 0.06 | 0.06 | 1.9 | 82.82
| Analytical results for K in OREAS 263b (Certified Value 0.238 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.2452 | 0.2448 | 0.21242 | 0.28 | 0.24 | 0.23 | 0.26 | 0.23 | 0.23 | 0.22 | 0.2368 | 0.344 | 0.23 | 0.3615 | 0.21 | 0.445 | 0.27 | 0.21 | 0.26894 |  |  | 1
| 1 | 2 | 0.2472 | 0.248 | 0.21393 | 0.29 | 0.24 | 0.23 | 0.26 | 0.22 | 0.23 | 0.22 | 0.23812 | 0.338 | 0.24 | 0.3614 | 0.22 | 0.427 | 0.26 | 0.21 | 0.27137 |  |  | #N/A
| 1 | 3 | 0.2504 | 0.24327 | 0.21615 | 0.29 | 0.24 | 0.22 | 0.26 | 0.23 | 0.22 | 0.23 | 0.23722 | 0.355 | 0.21 | 0.3725 | 0.22 | 0.443 | 0.27 | 0.22 | 0.25607 |  |  | 16
| 1 | 4 | 0.2449 | 0.24545 | 0.202 | 0.3 | 0.24 | 0.24 | 0.26 | 0.22 | 0.23 | 0.21 | 0.24022 | 0.355 | 0.21 | 0.3649 | 0.22 | 0.444 | 0.26 | 0.22 | 0.25147 |  |  | 0.23826
| 1 | 5 | 0.2462 | 0.24612 | 0.21044 | 0.3 | 0.24 | 0.22 | 0.26 | 0.22 | 0.23 | 0.22 | 0.23311 | 0.358 | 0.23 | 0.3658 | 0.22 | 0.39 | 0.26 | 0.21 | 0.2504 |  |  | 94
| 1 | 6 | 0.2464 | 0.25171 | 0.21246 | 0.26 | 0.24 | 0.22 | 0.26 | 0.22 | 0.23 | 0.21 | 0.23458 | 0.353 | 0.23 | 0.3451 | 0.23 | 0.489 | 0.27 | 0.22 | 0.26677
| Mean |  | 0.24672 | 0.24656 | 0.21123 | 0.28667 | 0.24 | 0.22667 | 0.26 | 0.22333 | 0.22833 | 0.21833 | 0.23668 | 0.3505 | 0.225 | 0.36187 | 0.22 | 0.43967 | 0.265 | 0.215 | 0.26083
| Median |  | 0.2463 | 0.24579 | 0.21244 | 0.29 | 0.24 | 0.225 | 0.26 | 0.22 | 0.23 | 0.22 | 0.23701 | 0.354 | 0.23 | 0.3632 | 0.22 | 0.4435 | 0.265 | 0.215 | 0.26142
| Std Dev. |  | 0.0019884 | 0.0029664 | 0.0049041 | 0.015055 | 0 | 0.008165 | 0 | 0.005164 | 0.0040825 | 0.0075277 | 0.0025328 | 0.0077653 | 0.012247 | 0.0091566 | 0.0063246 | 0.031998 | 0.0054772 | 0.0054772 | 0.0092869
| Rel.Std.Dev. |  | 0.8059% | 1.203% | 2.322% | 5.252% | 0% | 3.602% | 0% | 2.312% | 1.788% | 3.448% | 1.07% | 2.215% | 5.443% | 2.53% | 2.875% | 7.278% | 2.067% | 2.548% | 3.56%
| PDM3 |  | 3.548% | 3.482% | -11.34% | 20.32% | 0.7292% | -4.867% | 9.123% | -6.266% | -4.167% | -8.364% | -0.6657% | 47.11% | -5.566% | 51.88% | -7.665% | 84.53% | 11.22% | -9.763% | 9.474%
| Z-Score (Absolute) |  | 0.23 | 0.22 | 0.97 | 1.57 | 0 | 0.45 | 0.67 | 0.56 | 0.39 | 0.73 | 0.11 | 3.73 | 0.51 | 4.11 | 0.67 | 6.73 | 0.84 | 0.84 | 0.7
| Analytical results for La in OREAS 263b (Certified Value 23.2 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 1
| 1 | 1 | 22.249 | 21.907 | 39.794 | 27.6 | 23.8 | 22.1 | 24.5 | 20.8 | 21.6 | 20.6 | 23.9545 | 27.3 | 22.1 | 28 | 17.5 | 34.5 | 19.6 | 29.2 | 20.346 |  |  | 1
| 1 | 2 | 23.115 | 22.3789 | 39.8088 | 26.9 | 24 | 22.4 | 24.3 | 20.3 | 22.8 | 21.5 | 24.458 | 27 | 22.6 | 28 | 20.1 | 33.5 | 19.7 | 28.4 | 20.11 |  |  | 25
| 1 | 3 | 22.374 | 22.456 | 39.6964 | 27.3 | 23.7 | 20.8 | 24.4 | 21.5 | 21.4 | 20.7 | 22.7197 | 27.9 | 20.4 | 29 | 19.1 | 32.4 | 19.7 | 29.6 | 18.701 |  |  | 16
| 1 | 4 | 22.175 | 22.0406 | 39.6496 | 28.1 | 24.2 | 22.9 | 24.5 | 20.1 | 22.4 | 19.9 | 22.9093 | 27.4 | 21.9 | 29 | 20.5 | 35.6 | 19.3 | 30.8 | 18.423 |  |  | 23.1701
| 1 | 5 | 22.729 | 22.177 | 39.676 | 27.5 | 23.8 | 21.3 | 24.8 | 20.3 | 21.7 | 20.9 | 23.6776 | 27.8 | 22.3 | 28 | 18.6 | 30.6 | 19.7 | 29.2 | 18.408 |  |  | 95
| 1 | 6 | 22.692 | 22.7187 | 39.6508 | 26.4 | 23.7 | 21.1 | 24.2 | 20 | 23.1 | 20 | 24.2296 | 27.7 | 20.6 | 27 | 21.2 | 35.7 | 19.9 | 28.2 | 19.402
| Mean |  | 22.5557 | 22.2797 | 39.7126 | 27.3 | 23.8667 | 21.7667 | 24.45 | 20.5 | 22.1667 | 20.6 | 23.6581 | 27.5167 | 21.65 | 28.1667 | 19.5 | 33.7167 | 19.65 | 29.2333 | 19.2317
| Median |  | 22.533 | 22.2779 | 39.6862 | 27.4 | 23.8 | 21.7 | 24.45 | 20.3 | 22.05 | 20.65 | 23.816 | 27.55 | 22 | 28 | 19.6 | 34 | 19.7 | 29.2 | 19.0515
| Std Dev. |  | 0.35593 | 0.29677 | 0.071092 | 0.58992 | 0.19664 | 0.82381 | 0.20736 | 0.56214 | 0.70048 | 0.5933 | 0.70651 | 0.34303 | 0.9225 | 0.75277 | 1.3579 | 1.9793 | 0.19748 | 0.9331 | 0.85513
| Rel.Std.Dev. |  | 1.578% | 1.332% | 0.179% | 2.161% | 0.8239% | 3.785% | 0.8481% | 2.742% | 3.16% | 2.88% | 2.986% | 1.247% | 4.261% | 2.673% | 6.964% | 5.87% | 1.005% | 3.192% | 4.446%
| PDM3 |  | -2.652% | -3.843% | 71.4% | 17.82% | 3.006% | -6.057% | 5.524% | -11.52% | -4.331% | -11.09% | 2.106% | 18.76% | -6.561% | 21.56% | -15.84% | 45.52% | -15.19% | 26.17% | -17%
| Z-Score (Absolute) |  | 0 | 0.09 | 5.63 | 1.56 | 0.43 | 0.26 | 0.62 | 0.67 | 0.13 | 0.64 | 0.36 | 1.63 | 0.3 | 1.84 | 1 | 3.66 | 0.95 | 2.19 | 1.09
| Analytical results for Li in OREAS 263b (Certified Value 8.03 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 3
| 1 | 1 | 8.5 | < 10 | 8.5884 | 10.8 | 7.9 | 7.6 | 8.9 | 7.9 | 7.9 | 7.6 | 8.6644 | 11.2 | 11.8 | < 10 | 7.83 | [6 empty] | 1
| 1 | 2 | 8.4 | < 10 | 8.2404 | 11.4 | 7.6 | 7.5 | 8.8 | 7.6 | 8.2 | 7.6 | 8.7652 | 11.1 | 11.6 | < 10 | 7.773 | [6 empty] | #N/A
| 1 | 3 | 8.6 | < 10 | 9.1881 | 10.8 | 7.8 | 7.3 | 8.7 | 7.9 | 7.8 | 7.9 | 7.5841 | 11.4 | 16.5 | < 10 | 7.589 | [6 empty] | 16
| 1 | 4 | 8.6 | < 10 | 8.472 | 11 | 7.9 | 7.8 | 8.8 | 7.6 | 8 | 7.2 | 7.5983 | 11.4 | 12 | < 10 | 6.972 | [6 empty] | 8.0317
| 1 | 5 | 9 | < 10 | 8.5432 | 11.2 | 8.3 | 7.4 | 8.8 | 7.8 | 8.2 | 7.7 | 8.2906 | 11.7 | 12.8 | < 10 | 6.914 | [6 empty] | 96
| 1 | 6 | 8.8 | < 10 | 8.4291 | 9.9 | 8.1 | 7.3 | 8.8 | 7.6 | 8 | 7.3 | 8.5829 | 11.8 | 12.1 | < 10 | 7.608
| Mean |  | 8.65 |  | 8.5769 | 10.85 | 7.9333 | 7.4833 | 8.8 | 7.7333 | 8.0167 | 7.55 | 8.2476 | 11.4333 | 12.8 |  | 7.4477
| Median |  | 8.6 |  | 8.5076 | 10.9 | 7.9 | 7.45 | 8.8 | 7.7 | 8 | 7.6 | 8.4368 | 11.4 | 12.05 |  | 7.5985
| Std Dev. |  | 0.21679 |  | 0.32277 | 0.52058 | 0.24221 | 0.19408 | 0.063246 | 0.15055 | 0.16021 | 0.25884 | 0.53249 | 0.27325 | 1.858 |  | 0.40218
| Rel.Std.Dev. |  | 2.506% |  | 3.763% | 4.798% | 3.053% | 2.593% | 0.7187% | 1.947% | 1.998% | 3.428% | 6.456% | 2.39% | 14.52% |  | 5.4%
| PDM3 |  | 7.699% |  | 6.788% | 35.09% | -1.224% | -6.827% | 9.566% | -3.714% | -0.1866% | -5.997% | 2.688% | 42.35% | 59.37% |  | -7.271%
| Z-Score (Absolute) |  | 0.75 | 3.57 | 0.66 | 3.36 | 0.1 | 0.63 | 0.93 | 0.34 | 0 | 0.55 | 0.27 | 4.05 | 5.67 | 3.57 | 0.67
| Analytical results for Lu in OREAS 263b (Indicative Value 0.065 ppm) | [22 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 3
| 1 | 1 | 0.061 | 0.055343 | < 0.1 | 0.07 | [17 empty] | 1
| 1 | 2 | 0.064 | 0.059432 | < 0.1 | 0.07 | [17 empty] | 5
| 1 | 3 | 0.064 | 0.065703 | < 0.1 | 0.07 | [17 empty] | 16
| 1 | 4 | 0.062 | 0.061199 | < 0.1 | 0.07 | [17 empty] | 0.064803
| 1 | 5 | 0.065 | 0.06445 | < 0.1 | 0.07 | [17 empty] | 11
| 1 | 6 | 0.063 | 0.061326 | < 0.1 | 0.07
| Mean |  | 0.063167 | 0.061242 |  | 0.07
| Median |  | 0.0635 | 0.061263 |  | 0.07
| Std Dev. |  | 0.001472 | 0.0036952 |  | 0
| Rel.Std.Dev. |  | 2.33% | 6.034% |  | 0%
| PDM3 |  | -2.525% | -5.495% |  | 8.02%
| Z-Score (Absolute) |  | 0.15 | 0.15 | 1.88 | 1.2
| Analytical results for Mg in OREAS 263b (Certified Value 0.182 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.18 | 0.18598 | 0.16419 | 0.23 | 0.2 | 0.17 | 0.19 | 0.17 | 0.17 | 0.15 | 0.19979 | 0.23 | 0.18 | 0.25 | 0.16 | 0.234 | 0.18 | 0.19 | 0.19078 |  |  | 1
| 1 | 2 | 0.19 | 0.18788 | 0.17308 | 0.25 | 0.19 | 0.17 | 0.19 | 0.17 | 0.18 | 0.15 | 0.20475 | 0.23 | 0.18 | 0.25 | 0.17 | 0.233 | 0.18 | 0.19 | 0.20081 |  |  | #N/A
| 1 | 3 | 0.19 | 0.18487 | 0.16581 | 0.22 | 0.19 | 0.17 | 0.19 | 0.18 | 0.17 | 0.16 | 0.18898 | 0.24 | 0.18 | 0.25 | 0.17 | 0.218 | 0.18 | 0.19 | 0.19172 |  |  | 16
| 1 | 4 | 0.18 | 0.18576 | 0.1563 | 0.24 | 0.2 | 0.18 | 0.18 | 0.17 | 0.17 | 0.15 | 0.1918 | 0.24 | 0.17 | 0.25 | 0.18 | 0.268 | 0.18 | 0.19 | 0.18063 |  |  | 0.1821
| 1 | 5 | 0.19 | 0.18802 | 0.16524 | 0.24 | 0.19 | 0.17 | 0.18 | 0.17 | 0.18 | 0.16 | 0.19752 | 0.24 | 0.18 | 0.25 | 0.17 | 0.244 | 0.18 | 0.19 | 0.18012 |  |  | 97
| 1 | 6 | 0.19 | 0.18659 | 0.16545 | 0.21 | 0.19 | 0.17 | 0.18 | 0.17 | 0.17 | 0.15 | 0.20168 | 0.24 | 0.18 | 0.24 | 0.18 | 0.279 | 0.18 | 0.19 | 0.19374
| Mean |  | 0.18667 | 0.18652 | 0.16501 | 0.23167 | 0.19333 | 0.17167 | 0.185 | 0.17167 | 0.17333 | 0.15333 | 0.19742 | 0.23667 | 0.17833 | 0.24833 | 0.17167 | 0.246 | 0.18 | 0.19 | 0.18963
| Median |  | 0.19 | 0.18629 | 0.16534 | 0.235 | 0.19 | 0.17 | 0.185 | 0.17 | 0.17 | 0.15 | 0.19866 | 0.24 | 0.18 | 0.25 | 0.17 | 0.239 | 0.18 | 0.19 | 0.19125
| Std Dev. |  | 0.005164 | 0.0012401 | 0.0053387 | 0.01472 | 0.005164 | 0.0040825 | 0.0054772 | 0.0040825 | 0.005164 | 0.005164 | 0.0060068 | 0.005164 | 0.0040825 | 0.0040825 | 0.0075277 | 0.023126 | 3.0405e-17 | 3.0405e-17 | 0.0079877
| Rel.Std.Dev. |  | 2.766% | 0.6648% | 3.235% | 6.354% | 2.671% | 2.378% | 2.961% | 2.378% | 2.979% | 3.368% | 3.043% | 2.182% | 2.289% | 1.644% | 4.385% | 9.401% | 1.689e-14% | 1.6e-14% | 4.212%
| PDM3 |  | 2.508% | 2.427% | -9.383% | 27.22% | 6.169% | -5.729% | 1.593% | -5.729% | -4.814% | -15.8% | 8.413% | 29.97% | -2.068% | 36.37% | -5.729% | 35.09% | -1.153% | 4.339% | 4.138%
| Z-Score (Absolute) |  | 0.01 | 0 | 1.1 | 2.31 | 0.35 | 0.76 | 0.08 | 0.76 | 0.67 | 1.7 | 0.56 | 2.56 | 0.42 | 3.16 | 0.76 | 3.04 | 0.33 | 0.18 | 0.16
| Analytical results for Mn in OREAS 263b (Certified Value 0.013 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0129 | 0.013239 | 0.01284 | 0.0142 | 0.013 | 0.0125 | 0.0135 | 0.0129 | 0.0127 | 0.0126 | 0.013417 | 0.0143 | 0.0142 | 0.0143 | 0.0123 | 0.0138 | 0.01327 | 0.014 | 0.013509 |  |  | 1
| 1 | 2 | 0.0129 | 0.012867 | 0.012987 | 0.0147 | 0.013 | 0.0125 | 0.0136 | 0.0127 | 0.013 | 0.0125 | 0.013645 | 0.0138 | 0.0144 | 0.0144 | 0.013 | 0.0139 | 0.01309 | 0.014 | 0.013508 |  |  | 27
| 1 | 3 | 0.0132 | 0.013521 | 0.012938 | 0.0137 | 0.0131 | 0.0122 | 0.0135 | 0.0131 | 0.0127 | 0.0131 | 0.013597 | 0.0145 | 0.0141 | 0.0145 | 0.0131 | 0.0135 | 0.01313 | 0.014 | 0.013369 |  |  | 16
| 1 | 4 | 0.0128 | 0.01339 | 0.0126 | 0.0145 | 0.0132 | 0.0129 | 0.0133 | 0.0127 | 0.0131 | 0.0123 | 0.013535 | 0.0142 | 0.0138 | 0.0146 | 0.0131 | 0.0135 | 0.01319 | 0.014 | 0.013056 |  |  | 0.013385
| 1 | 5 | 0.0129 | 0.013579 | 0.012891 | 0.0142 | 0.0134 | 0.0125 | 0.0134 | 0.013 | 0.0131 | 0.0127 | 0.013895 | 0.0141 | 0.0139 | 0.0144 | 0.0129 | 0.0142 | 0.01294 | 0.014 | 0.013073 |  |  | 98
| 1 | 6 | 0.0131 | 0.013621 | 0.012627 | 0.0135 | 0.0132 | 0.0125 | 0.0134 | 0.0129 | 0.0129 | 0.0122 | 0.014004 | 0.0143 | 0.014 | 0.014 | 0.0131 | 0.014 | 0.01319 | 0.014 | 0.013411
| Mean |  | 0.012967 | 0.01337 | 0.012814 | 0.014133 | 0.01315 | 0.012517 | 0.01345 | 0.012883 | 0.012917 | 0.012567 | 0.013682 | 0.0142 | 0.014067 | 0.014367 | 0.012917 | 0.013817 | 0.013135 | 0.014 | 0.013321
| Median |  | 0.0129 | 0.013455 | 0.012865 | 0.0142 | 0.01315 | 0.0125 | 0.01345 | 0.0129 | 0.01295 | 0.01255 | 0.013621 | 0.01425 | 0.01405 | 0.0144 | 0.01305 | 0.01385 | 0.01316 | 0.014 | 0.01339
| Std Dev. |  | 0.00015055 | 0.00028274 | 0.00016294 | 0.00045898 | 0.00015166 | 0.00022286 | 0.00010488 | 0.00016021 | 0.00018348 | 0.00032042 | 0.00022327 | 0.00023664 | 0.00021602 | 0.00020656 | 0.00031252 | 0.00027869 | 0.00011345 | 0 | 0.00020596
| Rel.Std.Dev. |  | 1.161% | 2.115% | 1.272% | 3.248% | 1.153% | 1.781% | 0.7798% | 1.244% | 1.421% | 2.55% | 1.632% | 1.667% | 1.536% | 1.438% | 2.419% | 2.017% | 0.8637% | 0% | 1.546%
| PDM3 |  | -3.127% | -0.1161% | -4.269% | 5.589% | -1.757% | -6.489% | 0.4843% | -3.749% | -3.5% | -6.115% | 2.219% | 6.087% | 5.091% | 7.333% | -3.5% | 3.224% | -1.869% | 4.593% | -0.4788%
| Z-Score (Absolute) |  | 0.55 | 0.07 | 0.78 | 1.25 | 0.26 | 1.24 | 0.2 | 0.67 | 0.62 | 1.16 | 0.56 | 1.35 | 1.15 | 1.61 | 0.62 | 0.76 | 0.29 | 1.05 | 0
| Analytical results for Mo in OREAS 263b (Certified Value 2.15 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 2.3 | 2.1123 | 2.3051 | 2.31 | 2.01 | 2.11 | 2.11 | 1.95 | 2.06 | 2.22 | 2.0055 | 2.3 | 2.1 | 2.01 | 2 | 2.11 | 1.7 | 6.1 | 2.4 | 3.062 |  | 1
| 1 | 2 | 2.2 | 2.0634 | 2.198 | 2.26 | 2.02 | 2.13 | 2.1 | 1.89 | 2.11 | 2.34 | 2.0939 | 2.2 | 2.2 | 2.02 | 2 | 2.14 | 1.9 | 6.6 | 2.4 | 2.877 |  | 28
| 1 | 3 | 2.3 | 2.1179 | 2.3744 | 2.23 | 2.16 | 2.01 | 2.12 | 2.25 | 2.1 | 2.24 | 2.1604 | 2.3 | 2.1 | 1.91 | 2 | 2.09 | 2 | 6.4 | 2.4 | 2.768 |  | 16
| 1 | 4 | 2.3 | 2.1066 | 2.22 | 2.34 | 2.03 | 2.14 | 2.12 | 1.98 | 2.1 | 2.23 | 2.0228 | 2.4 | 2.2 | 2.24 | 2 | 2.17 | 2 | 6.7 | 2.4 | 2.649 |  | 2.155
| 1 | 5 | 2.3 | 2.192 | 2.2365 | 2.27 | 2.2 | 2.24 | 2.15 | 1.99 | 2.02 | 2.31 | 2.0807 | 2.2 | 2.2 | 2.01 | 2 | 2.13 | 2.3 | 5.9 | 2.4 | 2.602 |  | 99
| 1 | 6 | 2.3 | 2.176 | 2.307 | 2.18 | 2 | 2.03 | 2.1 | 2.01 | 2.13 | 2.2 | 2.1622 | 2.3 | 2.2 | 2.05 | 2 | 2.22 | 1.8 | 7.1 | 2.4 | 3.117
| Mean |  | 2.2833 | 2.128 | 2.2735 | 2.265 | 2.07 | 2.11 | 2.1167 | 2.0117 | 2.0867 | 2.2567 | 2.0876 | 2.2833 | 2.1667 | 2.04 | 2 | 2.1433 | 1.95 | 6.4667 | 2.4 | 2.8458
| Median |  | 2.3 | 2.1151 | 2.2708 | 2.265 | 2.025 | 2.12 | 2.115 | 1.985 | 2.1 | 2.235 | 2.0873 | 2.3 | 2.2 | 2.015 | 2 | 2.135 | 1.95 | 6.5 | 2.4 | 2.8225
| Std Dev. |  | 0.040825 | 0.047705 | 0.066677 | 0.056833 | 0.086718 | 0.083187 | 0.018619 | 0.124 | 0.039833 | 0.055377 | 0.066172 | 0.075277 | 0.05164 | 0.10881 | 0 | 0.046332 | 0.20736 | 0.43205 | 0 | 0.21236
| Rel.Std.Dev. |  | 1.788% | 2.242% | 2.933% | 2.509% | 4.189% | 3.942% | 0.8796% | 6.164% | 1.909% | 2.454% | 3.17% | 3.297% | 2.383% | 5.334% | 0% | 2.162% | 10.63% | 6.681% | 0% | 7.462%
| PDM3 |  | 5.956% | -1.252% | 5.5% | 5.105% | -3.944% | -2.088% | -1.778% | -6.651% | -3.17% | 4.718% | -3.128% | 5.956% | 0.542% | -5.336% | -7.192% | -0.5407% | -9.512% | 200.1% | 11.37% | 32.06%
| Z-Score (Absolute) |  | 0.83 | 0.09 | 0.77 | 0.72 | 0.44 | 0.2 | 0.16 | 0.78 | 0.34 | 0.67 | 0.33 | 0.83 | 0.14 | 0.61 | NA | 0 | 1.15 | 25.72 | 1.53 | 4.18
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 263b (Certified Value 0.019 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.02 | 0.017777 | 0.018971 | 0.024 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.017431 | 0.02 | 0.018 | 0.02 | 0.016 | 0.0176 | 0.01 | 0.02 | 0.0005192 |  |  | 1
| 1 | 2 | 0.02 | 0.018217 | 0.020445 | 0.026 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.017629 | 0.02 | 0.018 | 0.02 | 0.017 | 0.0154 | 0.01 | 0.02 | 0.023237 |  |  | #N/A
| 1 | 3 | 0.02 | 0.017831 | 0.019845 | 0.025 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.017913 | 0.02 | 0.017 | 0.02 | 0.017 | 0.0202 | 0.01 | 0.02 | 0.020075 |  |  | 16
| 1 | 4 | 0.02 | 0.018013 | 0.0179 | 0.025 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | < 0.01 | 0.017307 | 0.02 | 0.017 | 0.02 | 0.017 | 0.0171 | 0.01 | 0.03 | 0.0047014 |  |  | 0.019376
| 1 | 5 | 0.02 | 0.018051 | 0.018998 | 0.027 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.017407 | 0.02 | 0.017 | 0.02 | 0.016 | 0.0149 | 0.01 | 0.02 | 0.0032574 |  |  | 100
| 1 | 6 | 0.02 | 0.018082 | 0.019366 | 0.022 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.017832 | 0.02 | 0.018 | 0.02 | 0.017 | 0.0182 | 0.01 | 0.02 | 0.016592
| Mean |  | 0.02 | 0.017995 | 0.019254 | 0.024833 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.017586 | 0.02 | 0.0175 | 0.02 | 0.016667 | 0.017233 | 0.01 | 0.021667 | 0.011397
| Median |  | 0.02 | 0.018032 | 0.019182 | 0.025 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.01753 | 0.02 | 0.0175 | 0.02 | 0.017 | 0.01735 | 0.01 | 0.02 | 0.010647
| Std Dev. |  | 0 | 0.00016417 | 0.00086695 | 0.0017224 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0002461 | 0 | 0.00054772 | 0 | 0.0005164 | 0.0019336 | 0 | 0.0040825 | 0.0097148
| Rel.Std.Dev. |  | 0% | 0.9123% | 4.503% | 6.936% | 0% | 0% | 0% | 0% | 0% | 0% | 1.399% | 0% | 3.13% | 0% | 3.098% | 11.22% | 0% | 18.84% | 85.24%
| PDM3 |  | 3.222% | -7.125% | -0.6277% | 28.17% | 3.222% | -48.39% | 3.222% | 3.222% | 3.222% | -48.39% | -9.235% | 3.222% | -9.681% | 3.222% | -13.98% | -11.06% | -48.39% | 11.82% | -41.18%
| Z-Score (Absolute) |  | 0.3 | 0.51 | 0 | 2.26 | 0.3 | 3.74 | 0.3 | 0.3 | 0.3 | 4.08 | 0.67 | 0.3 | 0.71 | 0.3 | 1.05 | 0.82 | 3.74 | 0.98 | 3.18
| Analytical results for Nb in OREAS 263b (Certified Value 0.13 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 2
| 1 | 1 | 0.12 | 0.11803 | 0.1 | 0.65 | 0.12 | 0.13 | 0.11 | 0.1 | 0.14 | < 0.2 | < 0.2 | 0.15 | [9 empty] | 1
| 1 | 2 | 0.12 | 0.11469 | 0.1 | 0.68 | 0.12 | 0.13 | 0.11 | 0.1 | 0.15 | < 0.2 | < 0.2 | 0.15 | [9 empty] | 29
| 1 | 3 | 0.12 | 0.11413 | 0.1 | 0.65 | 0.11 | 0.13 | 0.11 | 0.11 | 0.18 | < 0.2 | < 0.2 | 0.17 | [9 empty] | 16
| 1 | 4 | 0.12 | 0.11221 | 0.1 | 0.68 | 0.12 | 0.13 | 0.1 | 0.11 | 0.15 | < 0.2 | < 0.2 | 0.18 | [9 empty] | 0.12678
| 1 | 5 | 0.13 | 0.12679 | 0.1 | 0.65 | 0.11 | 0.12 | 0.1 | 0.1 | 0.25 | < 0.2 | < 0.2 | 0.22 | [9 empty] | 101
| 1 | 6 | 0.12 | 0.12145 | 0.1 | 0.65 | 0.12 | 0.14 | 0.11 | 0.1 | 0.15 | < 0.2 | < 0.2 | 0.17
| Mean |  | 0.12167 | 0.11788 | 0.1 | 0.66 | 0.11667 | 0.13 | 0.10667 | 0.10333 | 0.17 |  |  | 0.17333
| Median |  | 0.12 | 0.11636 | 0.1 | 0.65 | 0.12 | 0.13 | 0.11 | 0.1 | 0.15 |  |  | 0.17
| Std Dev. |  | 0.0040825 | 0.0054512 | 1.5202e-17 | 0.015492 | 0.005164 | 0.0063246 | 0.005164 | 0.005164 | 0.041473 |  |  | 0.02582
| Rel.Std.Dev. |  | 3.355% | 4.624% | 1.52e-14% | 2.347% | 4.426% | 4.865% | 4.841% | 4.997% | 24.4% |  |  | 14.9%
| PDM3 |  | -4.031% | -7.014% | -21.12% | 420.6% | -7.975% | 2.542% | -15.86% | -18.49% | 34.09% |  |  | 36.72%
| Z-Score (Absolute) |  | 0.18 | 0 | NA | 25.12 | 0.06 | 0.56 | 0.52 | 0.67 | 2.42 | 0.83 | 0.83 | 2.57
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 263b (Indicative Value 21.4 ppm) | [22 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 1
| 1 | 1 | 20.568 | 19.6251 | 35.5608 | 23.1 | 21.6 | [16 empty] | 1
| 1 | 2 | 19.727 | 20.2704 | 35.6304 | 23.2 | 22.9 | [16 empty] | 6
| 1 | 3 | 20.342 | 20.1113 | 35.6176 | 23.2 | 21.3 | [16 empty] | 16
| 1 | 4 | 19.821 | 20.046 | 35.6824 | 23.1 | 22.6 | [16 empty] | 21.4051
| 1 | 5 | 20.582 | 20.855 | 35.578 | 23.1 | 22.3 | [16 empty] | 12
| 1 | 6 | 20.46 | 20.0155 | 35.572 | 22.3 | 22.6
| Mean |  | 20.25 | 20.1539 | 35.6069 | 23 | 22.2167
| Median |  | 20.401 | 20.0786 | 35.5978 | 23.1 | 22.45
| Std Dev. |  | 0.37986 | 0.40432 | 0.045925 | 0.34641 | 0.63061
| Rel.Std.Dev. |  | 1.876% | 2.006% | 0.129% | 1.506% | 2.838%
| PDM3 |  | -5.397% | -5.846% | 66.35% | 7.451% | 3.791%
| Z-Score (Absolute) |  | 0.67 | 0.71 | 4.59 | 0.27 | 0
| Analytical results for Ni in OREAS 263b (Certified Value 37.9 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 37 | 36.0152 | 38.087 | 37.45 | 39.2 | 39.1 | 38.4 | 38.2 | 37.3 | 38.4 | 36.7 | 37.0071 | 39 | 36.9 | 39 | 36 | 38.8 | 34.72 | 40 | 38.067 |  | 1
| 1 | 2 | 38.3 | 34.9234 | 38.786 | 33.89 | 39.8 | 38.4 | 41 | 37.4 | 36.2 | 39.5 | 37.2 | 36.9463 | 39 | 38.3 | 38 | 37.3 | 37.8 | 35.05 | 40 | 37.238 |  | 14
| 1 | 3 | 38.2 | 37.1101 | 39.324 | 37.74 | 38.9 | 39.5 | 38.2 | 38.1 | 37.3 | 38.8 | 36.2 | 38.0612 | 39 | 36.6 | 39 | 36.8 | 36.7 | 34.38 | 40 | 37.059 |  | 16
| 1 | 4 | 37.8 | 36.216 | 36 | 36.23 | 39.6 | 39.4 | 41.7 | 38.3 | 37.4 | 38.8 | 37.4 | 36.1982 | 39 | 38.7 | 40 | 37.9 | 37.7 | 34.85 | 40 | 36.208 |  | 37.9302
| 1 | 5 | 39.2 | 37.0123 | 38.485 | 37.3 | 39.8 | 39.4 | 40.5 | 37.8 | 36.6 | 37.7 | 38.2 | 36.3522 | 39 | 36.4 | 39 | 37.4 | 40.1 | 35.1 | 40 | 36.456 |  | 102
| 1 | 6 | 38.6 | 37.1063 | 38.156 | 37.69 | 37.6 | 38.1 | 38.7 | 38.1 | 36.4 | 39.1 | 35.8 | 38.4492 | 39 | 35.1 | 38 | 38.1 | 38.9 | 36.75 | 40 | 37.516
| Mean |  | 38.1833 | 36.3972 | 38.1397 | 36.7167 | 39.15 | 38.9833 | 39.75 | 37.9833 | 36.8667 | 38.7167 | 36.9167 | 37.169 | 39 | 37 | 38.8333 | 37.25 | 38.3333 | 35.1417 | 40 | 37.0907
| Median |  | 38.25 | 36.6142 | 38.3205 | 37.375 | 39.4 | 39.25 | 39.6 | 38.1 | 36.95 | 38.8 | 36.95 | 36.9767 | 39 | 36.75 | 39 | 37.35 | 38.3 | 34.95 | 40 | 37.1485
| Std Dev. |  | 0.74409 | 0.8649 | 1.142 | 1.4899 | 0.83845 | 0.59133 | 1.5003 | 0.33116 | 0.52789 | 0.61779 | 0.86814 | 0.90779 | 0 | 1.3206 | 0.75277 | 0.76616 | 1.1843 | 0.82944 | 0 | 0.68392
| Rel.Std.Dev. |  | 1.949% | 2.376% | 2.994% | 4.058% | 2.142% | 1.517% | 3.774% | 0.8719% | 1.432% | 1.596% | 2.352% | 2.442% | 0% | 3.569% | 1.938% | 2.057% | 3.09% | 2.36% | 0% | 1.844%
| PDM3 |  | 0.6675% | -4.042% | 0.5524% | -3.199% | 3.216% | 2.777% | 4.798% | 0.1402% | -2.804% | 2.074% | -2.672% | -2.007% | 2.821% | -2.452% | 2.381% | -1.793% | 1.063% | -7.352% | 5.457% | -2.213%
| Z-Score (Absolute) |  | 0.09 | 1.18 | 0.06 | 0.95 | 0.77 | 0.65 | 1.19 | 0.06 | 0.84 | 0.46 | 0.81 | 0.63 | 0.66 | 0.75 | 0.55 | 0.57 | 0.19 | 2.06 | 1.37 | 0.69
| Analytical results for P in OREAS 263b (Certified Value 0.023 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0225 | 0.023136 | 0.020151 | 0.026 | 0.0235 | 0.023 | 0.024 | 0.023 | 0.024 | 0.022 | 0.022307 | 0.0257 | 0.023 | 0.0256 | 0.019 | 0.0261 | 0.02 | 0.022 | 0.025754 |  |  | 1
| 1 | 2 | 0.0226 | 0.023509 | 0.020836 | 0.026 | 0.0234 | 0.023 | 0.024 | 0.022 | 0.024 | 0.022 | 0.022553 | 0.0251 | 0.023 | 0.0273 | 0.02 | 0.0261 | 0.02 | 0.022 | 0.025372 |  |  | #N/A
| 1 | 3 | 0.0231 | 0.023723 | 0.020906 | 0.026 | 0.0232 | 0.022 | 0.024 | 0.023 | 0.023 | 0.022 | 0.021481 | 0.026 | 0.023 | 0.0242 | 0.021 | 0.0263 | 0.02 | 0.022 | 0.024357 |  |  | 16
| 1 | 4 | 0.0224 | 0.023465 | 0.0194 | 0.026 | 0.0235 | 0.024 | 0.024 | 0.021 | 0.023 | 0.021 | 0.021415 | 0.0255 | 0.022 | 0.0256 | 0.021 | 0.0262 | 0.02 | 0.022 | 0.024015 |  |  | 0.023103
| 1 | 5 | 0.0226 | 0.023533 | 0.020874 | 0.026 | 0.0236 | 0.023 | 0.023 | 0.022 | 0.024 | 0.022 | 0.021908 | 0.0255 | 0.023 | 0.0247 | 0.02 | 0.0259 | 0.02 | 0.022 | 0.024101 |  |  | 103
| 1 | 6 | 0.0228 | 0.023892 | 0.020301 | 0.026 | 0.0244 | 0.023 | 0.023 | 0.022 | 0.023 | 0.021 | 0.022399 | 0.0258 | 0.022 | 0.0244 | 0.021 | 0.0266 | 0.02 | 0.022 | 0.025078
| Mean |  | 0.022667 | 0.023543 | 0.020411 | 0.026 | 0.0236 | 0.023 | 0.023667 | 0.022167 | 0.0235 | 0.021667 | 0.02201 | 0.0256 | 0.022667 | 0.0253 | 0.020333 | 0.0262 | 0.02 | 0.022 | 0.02478
| Median |  | 0.0226 | 0.023521 | 0.020569 | 0.026 | 0.0235 | 0.023 | 0.024 | 0.022 | 0.0235 | 0.022 | 0.022107 | 0.0256 | 0.023 | 0.02515 | 0.0205 | 0.02615 | 0.02 | 0.022 | 0.024718
| Std Dev. |  | 0.00025033 | 0.00025573 | 0.00059035 | 0 | 0.00041473 | 0.00063246 | 0.0005164 | 0.00075277 | 0.00054772 | 0.0005164 | 0.00048549 | 0.00030984 | 0.0005164 | 0.0011454 | 0.0008165 | 0.00023664 | 0 | 0 | 0.00072287
| Rel.Std.Dev. |  | 1.104% | 1.086% | 2.892% | 0% | 1.757% | 2.75% | 2.182% | 3.396% | 2.331% | 2.383% | 2.206% | 1.21% | 2.278% | 4.527% | 4.016% | 0.9032% | 0% | 0% | 2.917%
| PDM3 |  | -1.887% | 1.905% | -11.65% | 12.54% | 2.153% | -0.4445% | 2.441% | -4.052% | 1.72% | -6.216% | -4.728% | 10.81% | -1.887% | 9.511% | -11.99% | 13.41% | -13.43% | -4.773% | 7.259%
| Z-Score (Absolute) |  | 0.22 | 0.37 | 1.75 | 2.02 | 0.4 | 0 | 0.45 | 0.56 | 0.34 | 0.9 | 0.67 | 1.75 | 0.22 | 1.55 | 1.8 | 2.16 | 2.02 | 0.67 | 1.2
| Analytical results for Pb in OREAS 263b (Certified Value 12.9 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 12 | 13.9993 | 12.7876 | 14.1 | 14 | 11.4 | 13.4 | 11.6 | 12 | 12.1 | 13.0222 | 11 | 14 | 11.98 | 14 | 11.3 | 15.8 | 16.8 | 12 | 13.527 |  | 1
| 1 | 2 | 12.1 | 12.9753 | 13.431 | 14.3 | 15 | 11.2 | 13.2 | 11.2 | 12 | 11.9 | 12.1609 | 11 | 13.8 | 12.31 | 14 | 12.07 | 15.7 | 16.9 | 12 | 13.528 |  | 16
| 1 | 3 | 12.2 | 14.0058 | 14.7159 | 13.8 | 14 | 11.1 | 12.7 | 11.3 | 11.6 | 12.8 | 11.9777 | 27 | 14.2 | 11.38 | 14 | 11.51 | 15.5 | 17.3 | 12 | 12.179 |  | 16
| 1 | 4 | 11.9 | 14.1009 | 13.622 | 14.4 | 14.8 | 11.9 | 12.9 | 11.4 | 11.8 | 14.9 | 11.5546 | 12 | 14 | 12.47 | 14 | 12.19 | 16.1 | 17 | 12 | 12.969 |  | 12.9305
| 1 | 5 | 12.3 | 13.6747 | 12.765 | 14.3 | 13.7 | 11.6 | 13.1 | 11.2 | 12.2 | 12 | 12.4415 | 14 | 14.1 | 11.57 | 14 | 11.42 | 15.2 | 17.3 | 12 | 13.492 |  | 104
| 1 | 6 | 12 | 13.4543 | 13.9212 | 14 | 14.6 | 11.5 | 12.8 | 11.4 | 11.9 | 12.1 | 12.6096 | 14 | 14.1 | 11.67 | 13 | 12.26 | 17.1 | 17.9 | 11 | 13.847
| Mean |  | 12.0833 | 13.7018 | 13.5405 | 14.15 | 14.35 | 11.45 | 13.0167 | 11.35 | 11.9167 | 12.6333 | 12.2944 | 14.8333 | 14.0333 | 11.8967 | 13.8333 | 11.7917 | 15.9 | 17.2 | 11.8333 | 13.257
| Median |  | 12.05 | 13.837 | 13.5265 | 14.2 | 14.3 | 11.45 | 13 | 11.35 | 11.95 | 12.1 | 12.3012 | 13 | 14.05 | 11.825 | 14 | 11.79 | 15.75 | 17.15 | 12 | 13.5095
| Std Dev. |  | 0.1472 | 0.43127 | 0.73665 | 0.22583 | 0.52058 | 0.2881 | 0.26394 | 0.15166 | 0.20412 | 1.1553 | 0.51305 | 6.1128 | 0.13663 | 0.43163 | 0.40825 | 0.42771 | 0.6603 | 0.4 | 0.40825 | 0.59912
| Rel.Std.Dev. |  | 1.218% | 3.148% | 5.44% | 1.596% | 3.628% | 2.516% | 2.028% | 1.336% | 1.713% | 9.145% | 4.173% | 41.21% | 0.9736% | 3.628% | 2.951% | 3.627% | 4.153% | 2.326% | 3.45% | 4.519%
| PDM3 |  | -6.551% | 5.965% | 4.717% | 9.431% | 10.98% | -11.45% | 0.6666% | -12.22% | -7.84% | -2.298% | -4.919% | 14.72% | 8.529% | -7.995% | 6.982% | -8.807% | 22.97% | 33.02% | -8.485% | 2.525%
| Z-Score (Absolute) |  | 0.57 | 0.42 | 0.32 | 0.69 | 0.82 | 0.96 | 0 | 1.02 | 0.67 | 0.23 | 0.44 | NA | 0.62 | 0.69 | NA | 0.75 | 1.77 | 2.56 | NA | 0.15
Results from laboratories 1.15, 1.18 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 263b (Indicative Value < 10 ppb) | [22 empty] | Indicative
Pd | Round | Replicate | Lab | [20 empty] | 1
| No | No | 1.14 | [20 empty] | ppb
|  |  | AR*OES/MS | [20 empty] | 1
|  |  | 01g | [20 empty] | 1
| 1 | 1 | < 10 | [20 empty] | 1
| 1 | 2 | < 10 | [20 empty] | 7
| 1 | 3 | < 10 | [20 empty] | 16
| 1 | 4 | < 10 | [20 empty] | < 10
| 1 | 5 | < 10 | [20 empty] | 13
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 263b (Certified Value 6.43 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 5.174 | 5.1685 | 9.2659 | 6.6 | 5.81 | [16 empty] | 1
| 1 | 2 | 5.394 | 5.3109 | 9.3296 | 6.3 | 5.82 | [16 empty] | 33
| 1 | 3 | 5.376 | 5.329 | 9.2453 | 6.4 | 5.81 | [16 empty] | 16
| 1 | 4 | 5.243 | 5.2923 | 9.3195 | 6.7 | 5.81 | [16 empty] | 6.4342
| 1 | 5 | 5.475 | 5.1986 | 9.2487 | 6.4 | 5.77 | [16 empty] | 105
| 1 | 6 | 5.527 | 5.2943 | 9.3135 | 6.2 | 5.9
| Mean |  | 5.3648 | 5.2656 | 9.2871 | 6.4333 | 5.82
| Median |  | 5.385 | 5.2933 | 9.2897 | 6.4 | 5.81
| Std Dev. |  | 0.13466 | 0.065606 | 0.038028 | 0.18619 | 0.042895
| Rel.Std.Dev. |  | 2.51% | 1.246% | 0.4095% | 2.894% | 0.737%
| PDM3 |  | -16.62% | -18.16% | 44.34% | -0.01305% | -9.545%
| Z-Score (Absolute) |  | 0.55 | 0.67 | 4.22 | 0.75 | 0
| Analytical results for Pt in OREAS 263b (Indicative Value 2.17 ppb) | [22 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 1.14 | 1.21 | 1.28 | [18 empty] | ppb
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | [18 empty] | 2
|  |  | 01g | 0.1g | 40g | [18 empty] | 2
| 1 | 1 | < 5 | 2 | < 10 | [18 empty] | 1
| 1 | 2 | < 5 | 2 | < 10 | [18 empty] | 1
| 1 | 3 | < 5 | 1 | < 10 | [18 empty] | 16
| 1 | 4 | < 5 | 2 | < 10 | [18 empty] | 2.1667
| 1 | 5 | < 5 | 4 | < 10 | [18 empty] | 10
| 1 | 6 | < 5 | 2 | < 10
| Mean |  |  | 2.1667
| Median |  |  | 2
| Std Dev. |  |  | 0.98319
| Rel.Std.Dev. |  |  | 45.38%
| PDM3 |  |  | -1.665e-13%
| Z-Score (Absolute) |  | 0 | 0.67 | 5.06
| Analytical results for Rb in OREAS 263b (Certified Value 20.9 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 2
| 1 | 1 | 21.47 | 22.1752 | 23.4 | 21.3 | 19.8 | 19.4 | 19.9 | 19.4 | 19.4 | 21.8635 | 29.3 | 30.6 | 28.9 | [8 empty] | 1
| 1 | 2 | 22 | 21.3603 | 23.3 | 21.5 | 20.3 | 19.6 | 18.9 | 20.7 | 20.9 | 22.0026 | 29.5 | 30 | 28.8 | [8 empty] | 34
| 1 | 3 | 21.91 | 21.9967 | 23.3 | 21 | 19.3 | 19.6 | 19.8 | 20.5 | 20.3 | 21.037 | 30.5 | 28.2 | 28.9 | [8 empty] | 16
| 1 | 4 | 21.39 | 22.0943 | 24 | 21.1 | 21.3 | 19.4 | 19.4 | 21 | 19.8 | 20.7523 | 30.6 | 35.2 | 29.8 | [8 empty] | 20.9468
| 1 | 5 | 22.09 | 22.6843 | 23.6 | 21.4 | 19.8 | 19.5 | 19.1 | 19.6 | 20.4 | 21.5127 | 31.1 | 33.8 | 28.8 | [8 empty] | 106
| 1 | 6 | 22.07 | 22.3619 | 21.9 | 21.4 | 20 | 19.6 | 20 | 20.4 | 19.4 | 22.1755 | 30 | 39.4 | 28.6
| Mean |  | 21.8217 | 22.1121 | 23.25 | 21.2833 | 20.0833 | 19.5167 | 19.5167 | 20.2667 | 20.0333 | 21.5573 | 30.1667 | 32.8667 | 28.9667
| Median |  | 21.955 | 22.1347 | 23.35 | 21.35 | 19.9 | 19.55 | 19.6 | 20.45 | 20.05 | 21.6881 | 30.25 | 32.2 | 28.85
| Std Dev. |  | 0.3109 | 0.44105 | 0.71204 | 0.19408 | 0.67946 | 0.098319 | 0.4535 | 0.6314 | 0.60222 | 0.56476 | 0.69186 | 4.102 | 0.42269
| Rel.Std.Dev. |  | 1.425% | 1.995% | 3.063% | 0.9119% | 3.383% | 0.5038% | 2.324% | 3.115% | 3.006% | 2.62% | 2.293% | 12.48% | 1.459%
| PDM3 |  | 4.177% | 5.563% | 11% | 1.607% | -4.122% | -6.827% | -6.827% | -3.247% | -4.361% | 2.915% | 44.02% | 56.91% | 38.29%
| Z-Score (Absolute) |  | 0.12 | 0.25 | 0.75 | 0.12 | 0.65 | 0.9 | 0.9 | 0.57 | 0.67 | 0 | 3.81 | 5 | 3.28
| Analytical results for Re in OREAS 263b (Certified Value < 0.001 ppm) | [22 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 2
| 1 | 3 | < 0.01 | < 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | < 0.001 | 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | < 0.001
| 1 | 5 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 107
| 1 | 6 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.05 | < 0.001
| Mean |  |  |  |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001
| Median |  |  |  |  | 0.001 | 0.001 | 0.001 | 0.001 | 0.001
| Std Dev. |  |  |  |  | 0 | 0 | 0 | 0 | 0
| Rel.Std.Dev. |  |  |  |  | 0% | 0% | 0% | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 6.61 | 0.67 | 0.67 | 0.27 | 0.13 | 0 | 0 | 0.13 | 38.98 | 38.98 | 0.67
| Analytical results for S in OREAS 263b (Certified Value 0.027 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 3
| 1 | 1 | 0.0257 | 0.026725 | 0.022147 | 0.023694 | 0.026 | 0.02 | 0.03 | 0.04 | 0.03 | 0.05 | 0.02 | 0.030738 | 0.028 | 0.03 | 0.02 | 0.0218 | 0.02 | 0.0213 | 0.03 | 0.035 | 0.0302 | 1
| 1 | 2 | 0.026 | 0.027225 | 0.023129 | 0.022934 | 0.028 | 0.02 | 0.03 | 0.03 | 0.02 | 0.05 | 0.02 | 0.029275 | 0.0268 | 0.03 | 0.02 | 0.0234 | 0.03 | 0.0196 | 0.04 | 0.035 | 0.0278 | 20
| 1 | 3 | 0.0266 | 0.027044 | 0.021653 | 0.023422 | 0.026 | 0.02 | 0.03 | 0.04 | 0.03 | 0.05 | 0.03 | 0.031448 | 0.028 | 0.03 | 0.02 | 0.024 | 0.03 | 0.0199 | 0.04 | 0.035 | 0.0271 | 16
| 1 | 4 | 0.0259 | 0.027265 | 0.0203 | 0.023818 | 0.027 | 0.02 | 0.03 | 0.04 | 0.03 | 0.05 | 0.02 | 0.028996 | 0.0278 | 0.03 | 0.02 | 0.0261 | 0.03 | 0.0221 | 0.04 | 0.035 | 0.028 | 0.027092
| 1 | 5 | 0.0263 | 0.028421 | 0.022685 | 0.024219 | 0.027 | 0.02 | 0.03 | 0.03 | 0.03 | 0.05 | 0.03 | 0.027694 | 0.0279 | 0.03 | 0.02 | 0.018 | 0.03 | 0.0185 | 0.03 | 0.035 | 0.0272 | 108
| 1 | 6 | 0.0266 | 0.02858 | 0.021271 | 0.024789 | 0.026 | 0.02 | 0.03 | 0.04 | 0.03 | 0.05 | 0.02 | 0.029348 | 0.0287 | 0.03 | 0.02 | 0.0194 | 0.03 | 0.023 | 0.03 | 0.035 | 0.0312
| Mean |  | 0.026183 | 0.027543 | 0.021864 | 0.023813 | 0.026667 | 0.02 | 0.03 | 0.036667 | 0.028333 | 0.05 | 0.023333 | 0.029583 | 0.027867 | 0.03 | 0.02 | 0.022117 | 0.028333 | 0.020733 | 0.035 | 0.035 | 0.028583
| Median |  | 0.02615 | 0.027245 | 0.0219 | 0.023756 | 0.0265 | 0.02 | 0.03 | 0.04 | 0.03 | 0.05 | 0.02 | 0.029312 | 0.02795 | 0.03 | 0.02 | 0.0226 | 0.03 | 0.0206 | 0.035 | 0.035 | 0.0279
| Std Dev. |  | 0.00037639 | 0.00076734 | 0.0010194 | 0.00064152 | 0.0008165 | 0 | 0 | 0.005164 | 0.0040825 | 7.6012e-18 | 0.005164 | 0.0013323 | 0.0006121 | 0 | 0 | 0.0030162 | 0.0040825 | 0.0016908 | 0.0054772 | 0 | 0.0017046
| Rel.Std.Dev. |  | 1.438% | 2.786% | 4.662% | 2.694% | 3.062% | 0% | 0% | 14.08% | 14.41% | 1.52e-14% | 22.13% | 4.504% | 2.197% | 0% | 0% | 13.64% | 14.41% | 8.155% | 15.65% | 0% | 5.964%
| PDM3 |  | -3.353% | 1.667% | -19.3% | -12.1% | -1.569% | -26.18% | 10.74% | 35.34% | 4.583% | 84.56% | -13.87% | 9.197% | 2.861% | 10.74% | -26.18% | -18.36% | 4.583% | -23.47% | 29.19% | 29.19% | 5.506%
| Z-Score (Absolute) |  | 0.28 | 0.05 | 1 | 0.67 | 0.2 | 1.31 | 0.35 | 1.46 | 0.08 | 3.68 | 0.75 | 0.29 | 0 | 0.35 | 1.31 | 0.96 | 0.08 | 1.19 | 1.19 | 1.19 | 0.12
| Analytical results for Sb in OREAS 263b (Certified Value 7.35 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 8.48 | 7.1051 | 6.1452 | 8.84 | 8.22 | 6.76 | 7.1 | 6.37 | 6.42 | 8.01 | 4.8127 | 7.77 | 9.68 | 7.71 | 8 | 6.92 | 7.61 | < 5 | 8.9 | 5.055 |  | 1
| 1 | 2 | 8.79 | 6.9908 | 6.3599 | 7.76 | 8.16 | 6.97 | 7.16 | 6.13 | 6.54 | 8.31 | 5.0794 | 7.48 | 9.66 | 7.52 | 7 | 7.1 | 7.33 | < 5 | 8.9 | 4.656 |  | 36
| 1 | 3 | 8.88 | 7.412 | 6.4385 | 8.64 | 8.43 | 6.87 | 7.06 | 6.42 | 6.77 | 8.2 | 5.0346 | 7.74 | 9.76 | 7.91 | 7 | 6.94 | 7.35 | < 5 | 9 | 4.295 |  | 16
| 1 | 4 | 8.53 | 7.0827 | 6.046 | 8.87 | 8.23 | 6.89 | 6.88 | 6.18 | 6.83 | 8.15 | 4.744 | 7.71 | 9.39 | 8.32 | 7 | 6.99 | 7.72 | < 5 | 8.9 | 6.148 |  | 7.3509
| 1 | 5 | 8.81 | 7.1801 | 6.1124 | 8.04 | 8.65 | 6.74 | 6.7 | 6.15 | 6.33 | 7.93 | 4.8996 | 7.73 | 9.64 | 7.4 | 8 | 6.97 | 7.49 | < 5 | 8.8 | 4.306 |  | 109
| 1 | 6 | 8.6 | 7.5062 | 6.2129 | 8.83 | 8.23 | 6.87 | 7.15 | 6.27 | 6.69 | 8.12 | 4.6662 | 7.64 | 9.86 | 7.67 | 8 | 7.17 | 7.44 | < 5 | 8.8 | 6.886
| Mean |  | 8.6817 | 7.2128 | 6.2192 | 8.4967 | 8.32 | 6.85 | 7.0083 | 6.2533 | 6.5967 | 8.12 | 4.8728 | 7.6783 | 9.665 | 7.755 | 7.5 | 7.015 | 7.49 |  | 8.8833 | 5.2243
| Median |  | 8.695 | 7.1426 | 6.1791 | 8.735 | 8.23 | 6.87 | 7.08 | 6.225 | 6.615 | 8.135 | 4.8561 | 7.72 | 9.67 | 7.69 | 7.5 | 6.98 | 7.465 |  | 8.9 | 4.8555
| Std Dev. |  | 0.16606 | 0.20231 | 0.15155 | 0.47752 | 0.1859 | 0.086023 | 0.18203 | 0.12078 | 0.19937 | 0.13535 | 0.16275 | 0.10647 | 0.15719 | 0.32673 | 0.54772 | 0.098539 | 0.15166 |  | 0.075277 | 1.0652
| Rel.Std.Dev. |  | 1.913% | 2.805% | 2.437% | 5.62% | 2.234% | 1.256% | 2.597% | 1.931% | 3.022% | 1.667% | 3.34% | 1.387% | 1.626% | 4.213% | 7.303% | 1.405% | 2.025% |  | 0.8474% | 20.39%
| PDM3 |  | 18.1% | -1.878% | -15.4% | 15.59% | 13.18% | -6.814% | -4.66% | -14.93% | -10.26% | 10.46% | -33.71% | 4.454% | 31.48% | 5.497% | 2.028% | -4.569% | 1.892% |  | 20.85% | -28.93%
| Z-Score (Absolute) |  | 1.03 | 0 | 0.7 | 0.9 | 0.78 | 0.25 | 0.14 | 0.67 | 0.43 | 0.64 | 1.64 | 0.33 | 1.72 | 0.38 | NA | 0.14 | 0.19 | 3.31 | 1.17 | 1.4
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 263b (Certified Value 2.33 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 2.2 | 2.3345 | 2.6 | < 5 | 2.1 | 2.4 | 2.5 | 2.4 | 2.2 | 2.5063 | 3 | 2.3 | 3 | 1.9 | 3 | 2.2 | 2 | 2.336 |  |  |  | 1
| 1 | 2 | 2.2 | 2.237 | 2.7 | < 5 | 2.2 | 2.4 | 2.4 | 2.5 | 2.2 | 2.5251 | 3 | 2.3 | 3 | 2.1 | 2 | 2.2 | 2 | 2.317 |  |  |  | 37
| 1 | 3 | 2.3 | 2.3739 | 2.6 | < 5 | 2.1 | 2.5 | 2.5 | 2.4 | 2.3 | 2.4542 | 3 | 2.3 | 3 | 2.1 | 2 | 2.2 | 2 | 2.209 |  |  |  | 16
| 1 | 4 | 2.2 | 2.2868 | 2.7 | < 5 | 2.3 | 2.5 | 2.5 | 2.6 | 2.2 | 2.5173 | 3 | 2.4 | 3 | 2.1 | 3 | 2.2 | 2 | 2.157 |  |  |  | 2.3265
| 1 | 5 | 2.2 | 2.3125 | 2.6 | < 5 | 2.1 | 2.5 | 2.5 | 2.4 | 2.3 | 2.4669 | 3 | 2.3 | 3 | 2 | 3 | 2.2 | 2 | 2.149 |  |  |  | 110
| 1 | 6 | 2.3 | 2.3234 | 2.5 | < 5 | 2.1 | 2.5 | 2.5 | 2.6 | 2.1 | 2.495 | 3 | 2.2 | 3 | 2.1 | 3 | 2.2 | 2 | 2.268
| Mean |  | 2.2333 | 2.3113 | 2.6167 |  | 2.15 | 2.4667 | 2.4833 | 2.4833 | 2.2167 | 2.4941 | 3 | 2.3 | 3 | 2.05 | 2.6667 | 2.2 | 2 | 2.2393
| Median |  | 2.2 | 2.3179 | 2.6 |  | 2.1 | 2.5 | 2.5 | 2.45 | 2.2 | 2.5006 | 3 | 2.3 | 3 | 2.1 | 3 | 2.2 | 2 | 2.2385
| Std Dev. |  | 0.05164 | 0.046314 | 0.075277 |  | 0.083666 | 0.05164 | 0.040825 | 0.098319 | 0.075277 | 0.028225 | 0 | 0.063246 | 0 | 0.083666 | 0.5164 | 4.8648e-16 | 0 | 0.080072
| Rel.Std.Dev. |  | 2.312% | 2.004% | 2.877% |  | 3.891% | 2.094% | 1.644% | 3.959% | 3.396% | 1.132% | 0% | 2.75% | 0% | 4.081% | 19.36% | 2.211e-14% | 0% | 3.576%
| PDM3 |  | -4.006% | -0.6526% | 12.47% |  | -7.587% | 6.024% | 6.74% | 6.74% | -4.722% | 7.204% | 28.95% | -1.14% | 28.95% | -11.89% | 14.62% | -5.438% | -14.03% | -3.748%
| Z-Score (Absolute) |  | 0.31 | 0.02 | 1.32 | 0.83 | 0.66 | 0.69 | 0.76 | 0.76 | 0.38 | 0.8 | NA | 0.02 | NA | 1.09 | NA | 0.45 | NA | 0.28
Results from laboratories 1.16, 1.18, 1.21 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 263b (Certified Value 0.45 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1.5 | 1.1304 | 0.5 | < 5 | 0.5 | 0.4 | 0.4 | 0.4 | 0.5 | < 1 | 0.39 | < 1 | 0.2 | 0.4 | 1.07 | 0.5 | 2.373 |  |  |  |  | 1
| 1 | 2 | < 1.5 | 0.94472 | 0.5 | < 5 | 0.6 | 0.4 | 0.4 | 0.5 | 0.6 | < 1 | 0.37 | < 1 | 0.4 | 0.5 | 1.14 | 0.5 | [5 empty] | 38
| 1 | 3 | < 1.5 | 1.0715 | 0.4 | < 5 | 0.4 | 0.5 | 0.4 | 0.3 | 0.6 | < 1 | 0.39 | < 1 | 0.4 | 0.6 | 0.7 | 0.5 | 0.147 |  |  |  |  | 16
| 1 | 4 | < 1.5 | 1.018 | 0.5 | < 5 | 0.5 | 0.4 | 0.4 | 0.6 | 0.5 | < 1 | 0.4 | < 1 | 0.5 | 0.4 | 1.09 | 0.5 | 0.632 |  |  |  |  | 0.45233
| 1 | 5 | < 1.5 | 0.95463 | 0.5 | < 5 | 0.5 | 0.3 | 0.4 | 0.6 | 0.5 | < 1 | 0.4 | < 1 | 0.5 | 0.5 | 1.04 | 0.5 | 0.346 |  |  |  |  | 111
| 1 | 6 | < 1.5 | 0.89063 | 0.4 | < 5 | 0.4 | 0.3 | 0.4 | 0.6 | 0.6 | < 1 | 0.39 | < 1 | 0.1 | 0.3 | 0.81 | 0.5
| Mean |  |  | 1.0017 | 0.46667 |  | 0.48333 | 0.38333 | 0.4 | 0.5 | 0.55 |  | 0.39 |  | 0.35 | 0.45 | 0.975 | 0.5 | 0.8745
| Median |  |  | 0.98631 | 0.5 |  | 0.5 | 0.4 | 0.4 | 0.55 | 0.55 |  | 0.39 |  | 0.4 | 0.45 | 1.055 | 0.5 | 0.489
| Std Dev. |  |  | 0.088894 | 0.05164 |  | 0.075277 | 0.075277 | 6.0809e-17 | 0.12649 | 0.054772 |  | 0.010954 |  | 0.16432 | 0.10488 | 0.17695 | 0 | 1.0186
| Rel.Std.Dev. |  |  | 8.875% | 11.07% |  | 15.57% | 19.64% | 1.52e-14% | 25.3% | 9.959% |  | 2.809% |  | 46.95% | 23.31% | 18.15% | 0% | 116.5%
| PDM3 |  |  | 121.4% | 3.169% |  | 6.853% | -15.25% | -11.57% | 10.54% | 21.59% |  | -13.78% |  | -22.62% | -0.5158% | 115.5% | 10.54% | 93.33%
| Z-Score (Absolute) |  | 2.25 | 4.51 | 0.3 | NA | 0.15 | 1.05 | 0.9 | 0 | 0.45 | 0 | 0.99 | 0 | 1.35 | 0.45 | 4.27 | 0 | 3.37
Results from laboratory 1.08 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 263b (Indicative Value 3.82 ppm) | [22 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 3.715 | 3.5808 | 6.409 | 4.5 | 3.33 | [16 empty] | 1
| 1 | 2 | 3.694 | 3.6072 | 6.3091 | 4.5 | 3.51 | [16 empty] | 5
| 1 | 3 | 3.776 | 3.6893 | 6.394 | 4 | 3.33 | [16 empty] | 16
| 1 | 4 | 3.696 | 3.6578 | 6.3997 | 4.7 | 3.49 | [16 empty] | 3.8198
| 1 | 5 | 3.81 | 3.5974 | 6.4572 | 4.8 | 3.31 | [16 empty] | 11
| 1 | 6 | 3.812 | 3.6808 | 6.3387 | 4.4 | 3.49
| Mean |  | 3.7505 | 3.6355 | 6.3846 | 4.4833 | 3.41
| Median |  | 3.7455 | 3.6325 | 6.3968 | 4.5 | 3.41
| Std Dev. |  | 0.05549 | 0.046238 | 0.052908 | 0.27869 | 0.095499
| Rel.Std.Dev. |  | 1.48% | 1.272% | 0.8287% | 6.216% | 2.801%
| PDM3 |  | -1.815% | -4.825% | 67.14% | 17.37% | -10.73%
| Z-Score (Absolute) |  | 0 | 0.23 | 5.22 | 1.45 | 0.67
| Analytical results for Sn in OREAS 263b (Certified Value 1.69 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | 1.30 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | [5 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.66 | 1.7806 | 1.6365 | 2.4 | 1.8 | 1.6 | 1.6 | 1.5 | 1.6 | 1.7 | 1.8803 | 2.6 | 2.6 | < 5 | 2 | 0.455 | [5 empty] | 1
| 1 | 2 | 1.77 | 1.7309 | 1.7133 | 2.36 | 1.8 | 1.7 | 1.6 | 1.4 | 1.7 | 1.8 | 1.9182 | 2.6 | 2.5 | < 5 | 2 | [6 empty] | 11
| 1 | 3 | 1.73 | 1.7121 | 1.7715 | 2.3 | 1.9 | 1.6 | 1.6 | 1.5 | 1.6 | 1.7 | 1.8296 | 2.5 | 2.4 | < 5 | 2 | 0.223 | [5 empty] | 16
| 1 | 4 | 1.76 | 1.7336 | 1.615 | 2.45 | 1.9 | 1.8 | 1.6 | 1.5 | 1.7 | 1.7 | 1.8002 | 2.5 | 2.9 | < 5 | 2 | [6 empty] | 1.6921
| 1 | 5 | 1.71 | 1.7734 | 1.5871 | 2.39 | 1.8 | 1.6 | 1.6 | 1.5 | 1.6 | 1.6 | 1.85 | 2.4 | 2.6 | < 5 | 2 | [6 empty] | 112
| 1 | 6 | 1.79 | 1.75 | 1.6226 | 2.19 | 1.8 | 1.6 | 1.6 | 1.5 | 1.7 | 1.7 | 1.9006 | 2.7 | 3.2 | < 5 | 2
| Mean |  | 1.7367 | 1.7468 | 1.6577 | 2.3483 | 1.8333 | 1.65 | 1.6 | 1.4833 | 1.65 | 1.7 | 1.8631 | 2.55 | 2.7 |  | 2 | 0.339
| Median |  | 1.745 | 1.7418 | 1.6296 | 2.375 | 1.8 | 1.6 | 1.6 | 1.5 | 1.65 | 1.7 | 1.8652 | 2.55 | 2.6 |  | 2 | 0.339
| Std Dev. |  | 0.047188 | 0.02643 | 0.070032 | 0.09196 | 0.05164 | 0.083666 | 2.4324e-16 | 0.040825 | 0.054772 | 0.063246 | 0.044729 | 0.10488 | 0.29665 |  | 0 | 0.16405
| Rel.Std.Dev. |  | 2.717% | 1.513% | 4.225% | 3.916% | 2.817% | 5.071% | 1.52e-14% | 2.752% | 3.32% | 3.72% | 2.401% | 4.113% | 10.99% |  | 0% | 48.39%
| PDM3 |  | 2.634% | 3.231% | -2.034% | 38.78% | 8.347% | -2.488% | -5.442% | -12.34% | -2.488% | 0.4674% | 10.11% | 50.7% | 59.57% |  | 18.2% | -79.97%
| Z-Score (Absolute) |  | 0 | 0.05 | 0.42 | 3.26 | 0.52 | 0.46 | 0.73 | 1.35 | 0.46 | 0.2 | 0.67 | 4.34 | 5.14 | 4.07 | NA | 7.45
Results from laboratory 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 263b (Certified Value 17.3 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 1
| 1 | 1 | 17.38 | 17.746 | 11.347 | 28.6 | 18.8 | 16.8 | 19.3 | 16 | 14.2 | 15.1 | 21.4634 | 21.8 | 17.8 | 28 | 12.5 | 14.9 | 22.9 | 15 | 19.565 |  |  | 1
| 1 | 2 | 17.22 | 17.1058 | 11.431 | 29.8 | 19.8 | 17.6 | 20.1 | 15.5 | 14.8 | 16.5 | 21.4948 | 21.9 | 17.8 | 28 | 13.4 | 14.6 | 24.6 | 14 | 20.368 |  |  | 12
| 1 | 3 | 17.31 | 17.4933 | 11.69 | 27.9 | 18.1 | 16.7 | 19.9 | 15.9 | 14 | 15.2 | 20.6788 | 22 | 15.9 | 29 | 13.1 | 13.5 | 22.1 | 15 | 17.431 |  |  | 16
| 1 | 4 | 17.43 | 17.0676 | 11 | 30.5 | 19.4 | 18.6 | 19.5 | 15.7 | 14.8 | 15 | 20.5965 | 21.5 | 15.3 | 28 | 13.6 | 16.5 | 24.1 | 16 | 15.951 |  |  | 17.3283
| 1 | 5 | 17.32 | 17.0163 | 11.212 | 29.8 | 17.9 | 16.6 | 19.2 | 15.8 | 14.4 | 16.1 | 20.9459 | 21.9 | 16.7 | 28 | 12.8 | 14 | 25 | 15 | 15.938 |  |  | 113
| 1 | 6 | 17.14 | 17.9617 | 11.513 | 27.3 | 18.5 | 16.6 | 19.7 | 15.9 | 14.7 | 15.2 | 21.8436 | 21.9 | 18 | 27 | 14.5 | 17.4 | 24.9 | 14 | 18.973
| Mean |  | 17.3 | 17.3984 | 11.3655 | 28.9833 | 18.75 | 17.15 | 19.6167 | 15.8 | 14.4833 | 15.5167 | 21.1705 | 21.8333 | 16.9167 | 28 | 13.3167 | 15.15 | 23.9333 | 14.8333 | 18.0377
| Median |  | 17.315 | 17.2995 | 11.389 | 29.2 | 18.65 | 16.75 | 19.6 | 15.85 | 14.55 | 15.2 | 21.2046 | 21.9 | 17.25 | 28 | 13.25 | 14.75 | 24.35 | 15 | 18.202
| Std Dev. |  | 0.10564 | 0.39704 | 0.24035 | 1.2481 | 0.73959 | 0.80436 | 0.34881 | 0.17889 | 0.33714 | 0.62423 | 0.50313 | 0.17512 | 1.1339 | 0.63246 | 0.70261 | 1.503 | 1.181 | 0.75277 | 1.8851
| Rel.Std.Dev. |  | 0.6106% | 2.282% | 2.115% | 4.306% | 3.945% | 4.69% | 1.778% | 1.132% | 2.328% | 4.023% | 2.377% | 0.8021% | 6.703% | 2.259% | 5.276% | 9.921% | 4.934% | 5.075% | 10.45%
| PDM3 |  | -0.1631% | 0.4049% | -34.41% | 67.26% | 8.205% | -1.029% | 13.21% | -8.819% | -16.42% | -10.45% | 22.17% | 26% | -2.375% | 61.59% | -23.15% | -12.57% | 38.12% | -14.4% | 4.094%
| Z-Score (Absolute) |  | 0 | 0.03 | 1.86 | 3.66 | 0.45 | 0.05 | 0.73 | 0.47 | 0.88 | 0.56 | 1.21 | 1.42 | 0.12 | NA | 1.25 | 0.67 | 2.08 | NA | 0.23
Results from laboratories 1.18 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 263b (Certified Value < 0.01 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 25
| 1 | 3 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 114
| 1 | 6 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01
| Mean |  |  |  | 0.14
| Median |  |  |  | 0.14
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 263b (Certified Value 0.37 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.294 | 0.33864 | 0.4852 | 0.4 | 0.335 | [16 empty] | 1
| 1 | 2 | 0.299 | 0.3545 | 0.47368 | 0.4 | 0.325 | [16 empty] | 26
| 1 | 3 | 0.305 | 0.34611 | 0.50152 | 0.4 | 0.335 | [16 empty] | 16
| 1 | 4 | 0.295 | 0.35456 | 0.53584 | 0.4 | 0.34 | [16 empty] | 0.37032
| 1 | 5 | 0.292 | 0.33588 | 0.49 | 0.4 | 0.335 | [16 empty] | 115
| 1 | 6 | 0.289 | 0.346 | 0.47368 | 0.3 | 0.33
| Mean |  | 0.29567 | 0.34595 | 0.49332 | 0.38333 | 0.33333
| Median |  | 0.2945 | 0.34605 | 0.4876 | 0.4 | 0.335
| Std Dev. |  | 0.0056451 | 0.0077712 | 0.023339 | 0.040825 | 0.005164
| Rel.Std.Dev. |  | 1.909% | 2.246% | 4.731% | 10.65% | 1.549%
| PDM3 |  | -20.16% | -6.582% | 33.21% | 3.514% | -9.988%
| Z-Score (Absolute) |  | 0.91 | 0 | 2.66 | 0.67 | 0.23
| Analytical results for Te in OREAS 263b (Certified Value 0.22 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | [5 empty] | 3
| 1 | 1 | 0.2 | 0.21153 | 0.23 | 0.43 | 0.24 | 0.22 | 0.2 | 0.2 | 0.23 | 0.23937 | 0.18 | 0.22 | 0.2 | 0.24 | 0.24 | 0.4 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.20509 | 0.23 | 0.44 | 0.24 | 0.23 | 0.19 | 0.21 | 0.23 | 0.20283 | 0.2 | 0.22 | 0.17 | 0.23 | 0.24 | 0.4 | [5 empty] | 27
| 1 | 3 | 0.2 | < 0.2 | 0.25 | 0.42 | 0.21 | 0.21 | 0.21 | 0.22 | 0.21 | 0.2205 | 0.22 | 0.22 | 0.19 | 0.23 | 0.23 | 0.4 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.22201 | 0.27 | 0.43 | 0.24 | 0.23 | 0.19 | 0.21 | 0.23 | 0.20578 | 0.31 | 0.21 | 0.23 | 0.24 | 0.23 | 0.4 | [5 empty] | 0.22054
| 1 | 5 | 0.2 | 0.20351 | 0.25 | 0.41 | 0.22 | 0.24 | 0.2 | 0.19 | 0.23 | 0.20744 | 0.23 | 0.23 | 0.27 | 0.24 | 0.25 | 0.4 | [5 empty] | 116
| 1 | 6 | 0.2 | 0.19896 | 0.22 | 0.42 | 0.2 | 0.22 | 0.21 | 0.19 | 0.2 | 0.23455 | 0.25 | 0.23 | 0.18 | 0.24 | 0.24 | 0.4
| Mean |  | 0.2 | 0.20822 | 0.24167 | 0.425 | 0.225 | 0.225 | 0.2 | 0.20333 | 0.22167 | 0.21841 | 0.23167 | 0.22167 | 0.20667 | 0.23667 | 0.23833 | 0.4
| Median |  | 0.2 | 0.20509 | 0.24 | 0.425 | 0.23 | 0.225 | 0.2 | 0.205 | 0.23 | 0.21397 | 0.225 | 0.22 | 0.195 | 0.24 | 0.24 | 0.4
| Std Dev. |  | 3.0405e-17 | 0.0089284 | 0.018348 | 0.010488 | 0.017607 | 0.010488 | 0.0089443 | 0.012111 | 0.013292 | 0.015662 | 0.04535 | 0.0075277 | 0.037238 | 0.005164 | 0.0075277 | 6.0809e-17
| Rel.Std.Dev. |  | 1.52e-14% | 4.288% | 7.592% | 2.468% | 7.825% | 4.661% | 4.472% | 5.956% | 5.996% | 7.171% | 19.58% | 3.396% | 18.02% | 2.182% | 3.158% | 1.52e-14%
| PDM3 |  | -9.312% | -5.584% | 9.582% | 92.71% | 2.024% | 2.024% | -9.312% | -7.8% | 0.5128% | -0.9628% | 5.047% | 0.5128% | -6.289% | 7.314% | 8.07% | 81.38%
| Z-Score (Absolute) |  | NA | 1.58 | 0.87 | 9.6 | 0.08 | 0.08 | 1.11 | 0.95 | 0.08 | 0.23 | 0.4 | 0.08 | 0.79 | 0.63 | 0.71 | NA
Results from laboratories 1.02 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 263b (Certified Value 10.7 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 2
| 1 | 1 | 9.82 | 10.1649 | 10.621 | 15.8812 | 12.2 | 11.6 | 10.5 | 11.6 | 10.1 | 10 | 10.3 | 11.0524 | 10.1 | 10.4 | 9.1 | 9.65 | 12.3 | 12.5 |  |  |  | 1
| 1 | 2 | 10 | 10.4061 | 10.558 | 15.8748 | 12.1 | 11.6 | 10.1 | 11.4 | 9.7 | 9.9 | 9.9 | 10.9056 | 10.7 | 10 | 10.2 | 9.47 | 11.9 | 12.9 |  |  |  | 28
| 1 | 3 | 9.84 | 10.3781 | 11.191 | 15.8644 | 11.9 | 11.6 | 9.9 | 11 | 9.8 | 9.5 | 10.6 | 10.6414 | 10.9 | 9.6 | 9.8 | 9.43 | 12 | 12.6 |  |  |  | 16
| 1 | 4 | 9.63 | 10.825 | 10 | 15.9444 | 12.3 | 11.7 | 10.7 | 11.3 | 9.9 | 9.6 | 10.2 | 10.6562 | 11 | 11.1 | 10.4 | 9.56 | 11.8 | 12.3 |  |  |  | 10.7121
| 1 | 5 | 10.04 | 10.229 | 10.884 | 15.9028 | 12.2 | 11.7 | 10.3 | 11.6 | 9.8 | 9.8 | 10.3 | 10.9209 | 11.1 | 9.8 | 9.7 | 9.23 | 12.1 | 12.6 |  |  |  | 117
| 1 | 6 | 9.86 | 10.5285 | 11.262 | 15.8768 | 11.8 | 11.7 | 10.5 | 11.1 | 9.9 | 9.7 | 10 | 11.1089 | 11.1 | 9.8 | 10.2 | 9.99 | 12.1 | 12.6
| Mean |  | 9.865 | 10.4219 | 10.7527 | 15.8907 | 12.0833 | 11.65 | 10.3333 | 11.3333 | 9.8667 | 9.75 | 10.2167 | 10.8809 | 10.8167 | 10.1167 | 9.9 | 9.555 | 12.0333 | 12.5833
| Median |  | 9.85 | 10.3921 | 10.7525 | 15.879 | 12.15 | 11.65 | 10.4 | 11.35 | 9.85 | 9.75 | 10.25 | 10.9132 | 10.95 | 9.9 | 10 | 9.515 | 12.05 | 12.6
| Std Dev. |  | 0.14612 | 0.23632 | 0.46702 | 0.029182 | 0.19408 | 0.054772 | 0.29439 | 0.25033 | 0.13663 | 0.18708 | 0.24833 | 0.1957 | 0.38166 | 0.55287 | 0.47329 | 0.25564 | 0.17512 | 0.19408
| Rel.Std.Dev. |  | 1.481% | 2.268% | 4.343% | 0.1836% | 1.606% | 0.4701% | 2.849% | 2.209% | 1.385% | 1.919% | 2.431% | 1.799% | 3.528% | 5.465% | 4.781% | 2.675% | 1.455% | 1.542%
| PDM3 |  | -7.908% | -2.709% | 0.3788% | 48.34% | 12.8% | 8.756% | -3.536% | 5.799% | -7.892% | -8.981% | -4.625% | 1.576% | 0.9763% | -5.558% | -7.581% | -10.8% | 12.33% | 17.47%
| Z-Score (Absolute) |  | 0.68 | 0.15 | 0.15 | 4.96 | 1.4 | 0.99 | 0.24 | 0.7 | 0.67 | 0.78 | 0.35 | 0.27 | 0.21 | 0.44 | 0.64 | 0.96 | 1.35 | 1.87
| Analytical results for Ti in OREAS 263b (Certified Value 0.015 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0121 | 0.014501 | 0.013154 | 0.02 | 0.012 | 0.016 | 0.017 | 0.014 | 0.015 | 0.015 | 0.020069 | 0.0194 | 0.008 | 0.0169 | 0.009 | 0.0246 | 0.03 | 0.01 | 0.02069 |  |  | 1
| 1 | 2 | 0.0121 | 0.014214 | 0.013227 | 0.02 | 0.014 | 0.016 | 0.017 | 0.015 | 0.015 | 0.015 | 0.020181 | 0.0188 | 0.009 | 0.0193 | 0.01 | 0.0239 | 0.03 | 0.01 | 0.020475 |  |  | 29
| 1 | 3 | 0.0125 | 0.014703 | 0.013372 | 0.02 | 0.013 | 0.015 | 0.018 | 0.015 | 0.014 | 0.016 | 0.018476 | 0.0194 | 0.008 | 0.0193 | 0.009 | 0.0245 | 0.03 | 0.01 | 0.017736 |  |  | 16
| 1 | 4 | 0.0118 | 0.014775 | 0.0127 | 0.02 | 0.012 | 0.016 | 0.018 | 0.014 | 0.015 | 0.015 | 0.018152 | 0.0189 | 0.008 | 0.0198 | 0.01 | 0.0248 | 0.03 | 0.01 | 0.016134 |  |  | 0.014892
| 1 | 5 | 0.012 | 0.014439 | 0.012914 | 0.02 | 0.013 | 0.016 | 0.017 | 0.014 | 0.015 | 0.015 | 0.020319 | 0.0196 | 0.008 | 0.0165 | 0.009 | 0.0225 | 0.03 | 0.01 | 0.01611 |  |  | 118
| 1 | 6 | 0.0121 | 0.014908 | 0.013048 | 0.02 | 0.013 | 0.015 | 0.017 | 0.014 | 0.015 | 0.015 | 0.02066 | 0.0206 | 0.008 | 0.0164 | 0.01 | 0.0274 | 0.03 | 0.01 | 0.019523
| Mean |  | 0.0121 | 0.01459 | 0.013069 | 0.02 | 0.012833 | 0.015667 | 0.017333 | 0.014333 | 0.014833 | 0.015167 | 0.019643 | 0.01945 | 0.0081667 | 0.018033 | 0.0095 | 0.024617 | 0.03 | 0.01 | 0.018445
| Median |  | 0.0121 | 0.014602 | 0.013101 | 0.02 | 0.013 | 0.016 | 0.017 | 0.014 | 0.015 | 0.015 | 0.020125 | 0.0194 | 0.008 | 0.0181 | 0.0095 | 0.02455 | 0.03 | 0.01 | 0.018629
| Std Dev. |  | 0.00022804 | 0.00025305 | 0.00023877 | 0 | 0.00075277 | 0.0005164 | 0.0005164 | 0.0005164 | 0.00040825 | 0.00040825 | 0.0010534 | 0.0006442 | 0.00040825 | 0.0015895 | 0.00054772 | 0.0015993 | 0 | 0 | 0.0020793
| Rel.Std.Dev. |  | 1.885% | 1.734% | 1.827% | 0% | 5.866% | 3.296% | 2.979% | 3.603% | 2.752% | 2.692% | 5.363% | 3.312% | 4.999% | 8.815% | 5.766% | 6.497% | 0% | 0% | 11.27%
| PDM3 |  | -18.75% | -2.028% | -12.24% | 34.3% | -13.82% | 5.202% | 16.39% | -3.751% | -0.3937% | 1.845% | 31.9% | 30.61% | -45.16% | 21.09% | -36.21% | 65.3% | 101.5% | -32.85% | 23.86%
| Z-Score (Absolute) |  | 0.67 | 0.13 | 0.46 | 1.06 | 0.51 | 0.11 | 0.48 | 0.18 | 0.07 | 0 | 0.98 | 0.94 | 1.54 | 0.63 | 1.25 | 2.08 | 3.26 | 1.14 | 0.72
| Analytical results for Tl in OREAS 263b (Certified Value 0.43 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 3
| 1 | 1 | 0.44 | 0.44212 | 0.45272 | 0.47 | 0.44 | 0.4 | 0.46 | 0.37 | 0.37 | 0.41 | 0.45134 | 0.411 | 0.49 | 0.43 | 0.41 | 0.54 | < 5 | 0.4 |  |  |  | 1
| 1 | 2 | 0.43 | 0.4261 | 0.46973 | 0.47 | 0.44 | 0.39 | 0.44 | 0.36 | 0.37 | 0.4 | 0.43937 | 0.404 | 0.48 | 0.41 | 0.44 | 0.55 | < 5 | 0.4 |  |  |  | 30
| 1 | 3 | 0.44 | 0.45812 | 0.47473 | 0.45 | 0.42 | 0.38 | 0.43 | 0.39 | 0.37 | 0.42 | 0.42239 | 0.418 | 0.48 | 0.39 | 0.41 | 0.54 | < 5 | 0.4 |  |  |  | 16
| 1 | 4 | 0.42 | 0.46317 | 0.441 | 0.48 | 0.43 | 0.41 | 0.45 | 0.38 | 0.37 | 0.42 | 0.42736 | 0.413 | 0.5 | 0.45 | 0.44 | 0.57 | < 5 | 0.4 |  |  |  | 0.42594
| 1 | 5 | 0.43 | 0.43128 | 0.46137 | 0.47 | 0.43 | 0.4 | 0.44 | 0.37 | 0.38 | 0.4 | 0.42437 | 0.415 | 0.49 | 0.38 | 0.4 | 0.55 | < 5 | 0.4 |  |  |  | 119
| 1 | 6 | 0.43 | 0.4308 | 0.42098 | 0.45 | 0.44 | 0.38 | 0.44 | 0.37 | 0.37 | 0.39 | 0.4515 | 0.415 | 0.52 | 0.41 | 0.43 | 0.62 | < 5 | 0.4
| Mean |  | 0.43167 | 0.44193 | 0.45342 | 0.465 | 0.43333 | 0.39333 | 0.44333 | 0.37333 | 0.37167 | 0.40667 | 0.43605 | 0.41267 | 0.49333 | 0.41167 | 0.42167 | 0.56167 |  | 0.4
| Median |  | 0.43 | 0.4367 | 0.45705 | 0.47 | 0.435 | 0.395 | 0.44 | 0.37 | 0.37 | 0.405 | 0.43336 | 0.414 | 0.49 | 0.41 | 0.42 | 0.55 |  | 0.4
| Std Dev. |  | 0.0075277 | 0.015497 | 0.019951 | 0.012247 | 0.008165 | 0.012111 | 0.010328 | 0.010328 | 0.0040825 | 0.012111 | 0.013281 | 0.0048442 | 0.015055 | 0.025626 | 0.017224 | 0.030605 |  | 6.0809e-17
| Rel.Std.Dev. |  | 1.744% | 3.507% | 4.4% | 2.634% | 1.884% | 3.079% | 2.33% | 2.766% | 1.098% | 2.978% | 3.046% | 1.174% | 3.052% | 6.225% | 4.085% | 5.449% |  | 1.52e-14%
| PDM3 |  | 1.345% | 3.755% | 6.453% | 9.171% | 1.736% | -7.655% | 4.084% | -12.35% | -12.74% | -4.525% | 2.375% | -3.116% | 15.82% | -3.351% | -1.003% | 31.87% |  | -6.09%
| Z-Score (Absolute) |  | 0.05 | 0.27 | 0.63 | 0.99 | 0 | 1.24 | 0.31 | 1.87 | 1.92 | 0.83 | 0.08 | 0.64 | 1.87 | 0.67 | 0.36 | 3.99 | 64.32 | NA
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 263b (Indicative Value 0.075 ppm) | [22 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 3
| 1 | 1 | 0.076 | 0.07295 | < 0.1 | 0.075 | [17 empty] | 1
| 1 | 2 | 0.076 | 0.074178 | < 0.1 | 0.07 | [17 empty] | 6
| 1 | 3 | 0.082 | 0.075673 | < 0.1 | 0.07 | [17 empty] | 16
| 1 | 4 | 0.08 | 0.077137 | < 0.1 | 0.075 | [17 empty] | 0.075304
| 1 | 5 | 0.078 | 0.072854 | < 0.1 | 0.07 | [17 empty] | 12
| 1 | 6 | 0.086 | 0.074678 | < 0.1 | 0.07
| Mean |  | 0.079667 | 0.074578 |  | 0.071667
| Median |  | 0.079 | 0.074428 |  | 0.07
| Std Dev. |  | 0.0038816 | 0.0016457 |  | 0.002582
| Rel.Std.Dev. |  | 4.872% | 2.207% |  | 3.603%
| PDM3 |  | 5.794% | -0.9636% |  | -4.83%
| Z-Score (Absolute) |  | 1.1 | 0.25 | 3.9 | 0.25
| Analytical results for U in OREAS 263b (Certified Value 1.31 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g |  |  |  | 3
| 1 | 1 | 1.21 | 1.2255 | 1.4532 | 2.2137 | 1.3 | 1.42 | 1.25 | 1.29 | 1.21 | 1.28 | 1.22 | 1.3878 | 1.46 | 1.4 | 1.34 | 1.2 | 1.22 | 1.45 |  |  |  | 1
| 1 | 2 | 1.22 | 1.2233 | 1.3794 | 2.2168 | 1.3 | 1.4 | 1.2 | 1.28 | 1.16 | 1.24 | 1.23 | 1.4387 | 1.43 | 1.37 | 1.32 | 1.3 | 1.22 | 1.45 |  |  |  | 32
| 1 | 3 | 1.26 | 1.2757 | 1.5606 | 2.2007 | 1.3 | 1.36 | 1.18 | 1.25 | 1.18 | 1.22 | 1.26 | 1.363 | 1.47 | 1.39 | 1.21 | 1.2 | 1.21 | 1.45 |  |  |  | 16
| 1 | 4 | 1.21 | 1.2523 | 1.4074 | 2.2487 | 1.3 | 1.42 | 1.24 | 1.28 | 1.16 | 1.22 | 1.21 | 1.3656 | 1.46 | 1.41 | 1.43 | 1.3 | 1.21 | 1.5 |  |  |  | 1.3053
| 1 | 5 | 1.27 | 1.2369 | 1.385 | 2.1825 | 1.4 | 1.38 | 1.21 | 1.31 | 1.18 | 1.28 | 1.23 | 1.448 | 1.45 | 1.37 | 1.23 | 1.2 | 1.19 | 1.45 |  |  |  | 120
| 1 | 6 | 1.19 | 1.1986 | 1.3599 | 2.1698 | 1.3 | 1.43 | 1.21 | 1.28 | 1.19 | 1.26 | 1.15 | 1.4455 | 1.44 | 1.39 | 1.26 | 1.3 | 1.32 | 1.45
| Mean |  | 1.2267 | 1.2354 | 1.4242 | 2.2054 | 1.3167 | 1.4017 | 1.215 | 1.2817 | 1.18 | 1.25 | 1.2167 | 1.4081 | 1.4517 | 1.3883 | 1.2983 | 1.25 | 1.2283 | 1.4583
| Median |  | 1.215 | 1.2312 | 1.3962 | 2.2072 | 1.3 | 1.41 | 1.21 | 1.28 | 1.18 | 1.25 | 1.225 | 1.4132 | 1.455 | 1.39 | 1.29 | 1.25 | 1.215 | 1.45
| Std Dev. |  | 0.031411 | 0.026529 | 0.074049 | 0.027891 | 0.040825 | 0.027142 | 0.025884 | 0.019408 | 0.018974 | 0.027568 | 0.036697 | 0.040457 | 0.01472 | 0.016021 | 0.081833 | 0.054772 | 0.046224 | 0.020412
| Rel.Std.Dev. |  | 2.561% | 2.147% | 5.199% | 1.265% | 3.101% | 1.936% | 2.13% | 1.514% | 1.608% | 2.205% | 3.016% | 2.873% | 1.014% | 1.154% | 6.303% | 4.382% | 3.763% | 1.4%
| PDM3 |  | -6.026% | -5.357% | 9.11% | 68.95% | 0.8693% | 7.381% | -6.919% | -1.812% | -9.601% | -4.238% | -6.792% | 7.873% | 11.21% | 6.36% | -0.5352% | -4.238% | -5.898% | 11.72%
| Z-Score (Absolute) |  | 0.58 | 0.5 | 1.22 | 8.32 | 0.24 | 1.02 | 0.68 | 0.08 | 1 | 0.36 | 0.67 | 1.07 | 1.47 | 0.89 | 0.08 | 0.36 | 0.56 | 1.53
| Analytical results for V in OREAS 263b (Certified Value 22 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 22 | 21.2458 | 21.4435 | 24 | 22 | 22 | 23 | 22 | 23 | 21 | 23.2791 | 26 | 23 | 27 | 19 | 25 | 23.1 | 20 | 22.554 |  |  | 1
| 1 | 2 | 22 | 20.851 | 21.5524 | 25 | 22 | 22 | 24 | 22 | 23 | 21 | 23.301 | 26 | 23 | 26 | 21 | 24 | 23.2 | 20 | 22.523 |  |  | 33
| 1 | 3 | 22 | 21.7509 | 21.692 | 23 | 21 | 21 | 24 | 23 | 22 | 22 | 21.7889 | 27 | 23 | 27 | 21 | 23 | 22.9 | 30 | 21.63 |  |  | 16
| 1 | 4 | 22 | 21.5096 | 20.5224 | 24 | 22 | 23 | 23 | 22 | 23 | 20 | 21.8496 | 26 | 22 | 27 | 21 | 27 | 22.9 | 20 | 21.222 |  |  | 22.0087
| 1 | 5 | 22 | 21.6821 | 21.015 | 24 | 22 | 22 | 23 | 22 | 23 | 22 | 22.3588 | 26 | 22 | 27 | 20 | 28 | 22.7 | 20 | 21.212 |  |  | 121
| 1 | 6 | 22 | 21.8019 | 20.858 | 23 | 21 | 22 | 23 | 22 | 23 | 20 | 22.8216 | 26 | 23 | 26 | 21 | 29 | 23.5 | 20 | 22.069
| Mean |  | 22 | 21.4736 | 21.1806 | 23.8333 | 21.6667 | 22 | 23.3333 | 22.1667 | 22.8333 | 21 | 22.5665 | 26.1667 | 22.6667 | 26.6667 | 20.5 | 26 | 23.05 | 21.6667 | 21.8683
| Median |  | 22 | 21.5959 | 21.2292 | 24 | 22 | 22 | 23 | 22 | 23 | 21 | 22.5902 | 26 | 23 | 27 | 21 | 26 | 23 | 20 | 21.8495
| Std Dev. |  | 0 | 0.36577 | 0.45467 | 0.75277 | 0.5164 | 0.63246 | 0.5164 | 0.40825 | 0.40825 | 0.89443 | 0.67438 | 0.40825 | 0.5164 | 0.5164 | 0.83666 | 2.3664 | 0.28107 | 4.0825 | 0.60734
| Rel.Std.Dev. |  | 0% | 1.703% | 2.147% | 3.158% | 2.383% | 2.875% | 2.213% | 1.842% | 1.788% | 4.259% | 2.988% | 1.56% | 2.278% | 1.936% | 4.081% | 9.102% | 1.219% | 18.84% | 2.777%
| PDM3 |  | -0.03946% | -2.431% | -3.763% | 8.291% | -1.554% | -0.03946% | 6.019% | 0.7178% | 3.747% | -4.583% | 2.535% | 18.89% | 2.99% | 21.16% | -6.855% | 18.14% | 4.731% | -1.554% | -0.6377%
| Z-Score (Absolute) |  | 0.16 | 0.67 | 0.96 | 1.62 | 0.49 | 0.16 | 1.14 | 0 | 0.65 | 1.14 | 0.39 | 3.89 | 0.49 | 4.38 | 1.62 | 3.73 | 0.86 | 0.49 | 0.29
| Analytical results for W in OREAS 263b (Certified Value 14.8 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g |  |  |  |  | 2
| 1 | 1 | 13.42 | 16.7753 | 15.269 | 13.8 | 15.07 | 15.9 | 15.35 | 14.25 | 14.2 | 15.35 | 14.3894 | 16.3 | 12.97 | 11 | 11.4 | 18.1 | 10.2 |  |  |  |  | 1
| 1 | 2 | 13.9 | 15.8918 | 16.8721 | 11.6 | 14.94 | 14.8 | 15.7 | 14.3 | 13.05 | 14.8 | 14.6412 | 16.3 | 13.58 | 10 | 13.5 | 17 | 11.5 |  |  |  |  | 34
| 1 | 3 | 13.89 | 15.5125 | 15.899 | 14.1 | 14.86 | 14.3 | 15.8 | 14 | 13.85 | 15.85 | 14.4063 | 16 | 13.47 | 12 | 11.9 | 16.1 | 9.5 |  |  |  |  | 16
| 1 | 4 | 12.91 | 16.4426 | 17.2476 | 13.9 | 14.97 | 15.3 | 15.15 | 13.45 | 13.55 | 15.65 | 14.5497 | 16.5 | 14.25 | 8 | 12.9 | 15.9 | 9.4 |  |  |  |  | 14.7639
| 1 | 5 | 14.42 | 17.3056 | 15.496 | 12 | 14.04 | 15.4 | 15.3 | 13.95 | 13.75 | 15.05 | 14.3967 | 16.1 | 13.38 | 11 | 11.8 | 16.2 | 11.4 |  |  |  |  | 122
| 1 | 6 | 14.38 | 16.2434 | 13.921 | 15.7 | 15.29 | 15.65 | 15.1 | 14.1 | 14.2 | 14.75 | 14.5802 | 16.6 | 12.14 | 11 | 12.8 | 17.7 | 11.5
| Mean |  | 13.82 | 16.3619 | 15.7841 | 13.5167 | 14.8617 | 15.225 | 15.4 | 14.0083 | 13.7667 | 15.2417 | 14.4939 | 16.3 | 13.2983 | 10.5 | 12.3833 | 16.8333 | 10.5833
| Median |  | 13.895 | 16.343 | 15.6975 | 13.85 | 14.955 | 15.35 | 15.325 | 14.05 | 13.8 | 15.2 | 14.478 | 16.3 | 13.425 | 11 | 12.35 | 16.6 | 10.8
| Std Dev. |  | 0.57845 | 0.63632 | 1.1965 | 1.5039 | 0.42893 | 0.58459 | 0.2881 | 0.30564 | 0.43436 | 0.45212 | 0.10982 | 0.22804 | 0.70295 | 1.3784 | 0.80353 | 0.91579 | 1.0068
| Rel.Std.Dev. |  | 4.186% | 3.889% | 7.581% | 11.13% | 2.886% | 3.84% | 1.871% | 2.182% | 3.155% | 2.966% | 0.7577% | 1.399% | 5.286% | 13.13% | 6.489% | 5.44% | 9.513%
| PDM3 |  | -6.394% | 10.82% | 6.91% | -8.448% | 0.6619% | 3.123% | 4.308% | -5.118% | -6.755% | 3.236% | -1.829% | 10.4% | -9.927% | -28.88% | -16.12% | 14.02% | -28.32%
| Z-Score (Absolute) |  | 0.57 | 1.13 | 0.74 | 0.78 | 0.12 | 0.37 | 0.48 | 0.45 | 0.61 | 0.38 | 0.12 | 1.08 | 0.92 | NA | 1.53 | 1.44 | 2.74
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 263b (Certified Value 6.2 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 3
| 1 | 1 | 6.06 | 6.1493 | 6.0452 | 6.54 | 5.99 | 6.02 | 6.38 | 5.99 | 5.91 | 5.58 | 6.8587 | 7.99 | 7.51 | 6.25 | 7.239 | [6 empty] | 1
| 1 | 2 | 6.18 | 6.0741 | 5.7224 | 6.81 | 6.06 | 6.2 | 6.37 | 5.7 | 5.91 | 6.02 | 6.9521 | 7.77 | 7.47 | 6.2 | 7.169 | [6 empty] | 35
| 1 | 3 | 6.18 | 6.2347 | 5.797 | 6.49 | 6.01 | 5.91 | 6.35 | 5.9 | 5.79 | 5.77 | 6.3787 | 8.11 | 7.03 | 6.35 | 6.782 | [6 empty] | 16
| 1 | 4 | 6.02 | 6.1928 | 5.3448 | 6.74 | 6.14 | 6.54 | 6.38 | 5.83 | 6.06 | 6.07 | 6.4029 | 8.11 | 8.32 | 6.35 | 6.503 | [6 empty] | 6.1998
| 1 | 5 | 6.14 | 6.3023 | 5.731 | 6.64 | 6.21 | 6.21 | 6.26 | 5.67 | 5.87 | 5.99 | 6.7084 | 8.13 | 7.07 | 6.15 | 6.591 | [6 empty] | 123
| 1 | 6 | 6.21 | 6.3452 | 5.676 | 6.33 | 6 | 5.83 | 6.35 | 5.78 | 5.91 | 5.72 | 6.9506 | 8.15 | 8.55 | 6.15 | 6.961
| Mean |  | 6.1317 | 6.2164 | 5.7194 | 6.5917 | 6.0683 | 6.1183 | 6.3483 | 5.8117 | 5.9083 | 5.8583 | 6.7086 | 8.0433 | 7.6583 | 6.2417 | 6.8742
| Median |  | 6.16 | 6.2137 | 5.7267 | 6.59 | 6.035 | 6.11 | 6.36 | 5.805 | 5.91 | 5.88 | 6.7836 | 8.11 | 7.49 | 6.225 | 6.8715
| Std Dev. |  | 0.075476 | 0.099601 | 0.22561 | 0.1752 | 0.088863 | 0.25655 | 0.04535 | 0.12123 | 0.087731 | 0.19631 | 0.26178 | 0.14514 | 0.63738 | 0.091742 | 0.30133
| Rel.Std.Dev. |  | 1.231% | 1.602% | 3.945% | 2.658% | 1.464% | 4.193% | 0.7144% | 2.086% | 1.485% | 3.351% | 3.902% | 1.805% | 8.323% | 1.47% | 4.384%
| PDM3 |  | -1.098% | 0.268% | -7.748% | 6.321% | -2.12% | -1.313% | 2.396% | -6.26% | -4.701% | -5.507% | 8.207% | 29.74% | 23.53% | 0.676% | 10.88%
| Z-Score (Absolute) |  | 0.16 | 0 | 0.94 | 0.71 | 0.28 | 0.18 | 0.25 | 0.76 | 0.58 | 0.67 | 0.93 | 3.44 | 2.72 | 0.05 | 1.24
| Analytical results for Yb in OREAS 263b (Indicative Value 0.53 ppm) | [22 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.434 | 0.45176 | 0.71284 | 0.5 | [17 empty] | 1
| 1 | 2 | 0.469 | 0.43828 | 0.6754 | 0.5 | [17 empty] | 7
| 1 | 3 | 0.494 | 0.46587 | 0.73456 | 0.5 | [17 empty] | 16
| 1 | 4 | 0.447 | 0.43677 | 0.77148 | 0.5 | [17 empty] | 0.52992
| 1 | 5 | 0.463 | 0.42052 | 0.7072 | 0.5 | [17 empty] | 13
| 1 | 6 | 0.46 | 0.45939 | 0.67696 | 0.5
| Mean |  | 0.46117 | 0.44543 | 0.71307 | 0.5
| Median |  | 0.4615 | 0.44502 | 0.71002 | 0.5
| Std Dev. |  | 0.02041 | 0.016729 | 0.036417 | 0
| Rel.Std.Dev. |  | 4.426% | 3.756% | 5.107% | 0%
| PDM3 |  | -12.97% | -15.94% | 34.56% | -5.646%
| Z-Score (Absolute) |  | 0.48 | 0.87 | 5.75 | 0.48
| Analytical results for Zn in OREAS 263b (Certified Value 30 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 2
| 1 | 1 | 28 | 29.1367 | 28.1399 | 32.7875 | 32.4 | 29 | 29 | 29 | 28 | 32 | 28 | 31.3631 | 25 | 35 | 27.1 | 34 | 26.4 | 30.7 | 45.6 | 29 | 34.539 | 1
| 1 | 2 | 29 | 27.7922 | 31.2181 | 31.6375 | 40 | 28 | 29 | 29 | 27 | 33 | 28 | 29.9875 | 24 | 34 | 27.8 | 34 | 27.5 | 30.3 | 42.9 | 29 | 31.679 | 15
| 1 | 3 | 28 | 29.2383 | 28.237 | 33.213 | 32.3 | 29 | 28 | 30 | 28 | 32 | 30 | 30.281 | 25 | 36 | 25 | 34 | 27.2 | 29.1 | 41.2 | 29 | 30.794 | 16
| 1 | 4 | 28 | 28.9343 | 27.8126 | 33.19 | 33.2 | 29 | 30 | 29 | 27 | 33 | 33 | 31.5074 | 25 | 34 | 26.9 | 35 | 27.7 | 32.2 | 42 | 30 | 31.407 | 29.9529
| 1 | 5 | 28 | 29.3122 | 30.375 | 34.6505 | 34.9 | 30 | 28 | 29 | 28 | 33 | 28 | 32.0894 | 25 | 35 | 26.6 | 34 | 27.2 | 31.6 | 41.5 | 30 | 30.388 | 124
| 1 | 6 | 28 | 29.6517 | 29.858 | 32.638 | 35.4 | 29 | 29 | 29 | 27 | 32 | 27 | 31.2723 | 25 | 36 | 26.3 | 33 | 27.9 | 32.2 | 46 | 29 | 32.253
| Mean |  | 28.1667 | 29.0109 | 29.2734 | 33.0194 | 34.7 | 29 | 28.8333 | 29.1667 | 27.5 | 32.5 | 29 | 31.0835 | 24.8333 | 35 | 26.6167 | 34 | 27.3167 | 31.0167 | 43.2 | 29.3333 | 31.8433
| Median |  | 28 | 29.1875 | 29.0475 | 32.9887 | 34.05 | 29 | 29 | 29 | 27.5 | 32.5 | 28 | 31.3177 | 25 | 35 | 26.75 | 34 | 27.35 | 31.15 | 42.45 | 29 | 31.543
| Std Dev. |  | 0.40825 | 0.64189 | 1.4021 | 0.98346 | 2.8955 | 0.63246 | 0.75277 | 0.40825 | 0.54772 | 0.54772 | 2.1909 | 0.7938 | 0.40825 | 0.89443 | 0.9411 | 0.63246 | 0.52694 | 1.2189 | 2.0986 | 0.5164 | 1.4745
| Rel.Std.Dev. |  | 1.449% | 2.213% | 4.79% | 2.978% | 8.344% | 2.181% | 2.611% | 1.4% | 1.992% | 1.685% | 7.555% | 2.554% | 1.644% | 2.556% | 3.536% | 1.86% | 1.929% | 3.93% | 4.858% | 1.76% | 4.631%
| PDM3 |  | -5.964% | -3.145% | -2.269% | 10.24% | 15.85% | -3.181% | -3.738% | -2.625% | -8.189% | 8.504% | -3.181% | 3.774% | -17.09% | 16.85% | -11.14% | 13.51% | -8.801% | 3.551% | 44.23% | -2.069% | 6.311%
| Z-Score (Absolute) |  | 0.41 | 0.1 | 0 | 1.4 | 2.02 | 0.1 | 0.16 | 0.04 | 0.66 | 1.2 | 0.1 | 0.67 | 1.65 | 2.13 | 0.99 | 1.76 | 0.73 | 0.65 | 5.19 | 0.02 | 0.96
| Analytical results for Zr in OREAS 263b (Certified Value 15.7 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 2
| 1 | 1 | 15.4 | 14.0165 | 17.1402 | 4.1 | 17.8 | 14.5 | 15.6 | 14.9 | 15.1 | 15.3 | 15.8473 | 11.8 | 14.9 | 16.1 | 6 | [6 empty] | 1
| 1 | 2 | 16 | 14 | 15.9552 | 1.8 | 17.5 | 15.3 | 15.5 | 14.9 | 15.3 | 16.1 | 15.3849 | 12.7 | 14.2 | 15.9 | 7 | [6 empty] | 16
| 1 | 3 | 15.9 | 14.0365 | 17.1206 | 3.9 | 17.3 | 15.1 | 15.8 | 15.6 | 15.6 | 15.9 | 15.8285 | 12.5 | 13.3 | 15.9 | 5 | [6 empty] | 16
| 1 | 4 | 15.5 | 14.3243 | 17.297 | 2.9 | 18.1 | 15 | 15.5 | 14.5 | 15.8 | 17.1 | 15.7293 | 12.4 | 15.2 | 15.5 | 6 | [6 empty] | 15.6975
| 1 | 5 | 15.9 | 14.5556 | 17.5126 | 1.7 | 17.2 | 14.9 | 14.4 | 14.2 | 14.9 | 16.6 | 15.3214 | 12.4 | 16 | 15.6 | 6 | [6 empty] | 125
| 1 | 6 | 15.8 | 14.3353 | 17.3754 | 5.6 | 17.3 | 14.7 | 15.4 | 14.9 | 16.1 | 16 | 15.6443 | 12.6 | 17 | 16.5 | 6
| Mean |  | 15.75 | 14.2114 | 17.0668 | 3.3333 | 17.5333 | 14.9167 | 15.3667 | 14.8333 | 15.4667 | 16.1667 | 15.626 | 12.4 | 15.1 | 15.9167 | 6
| Median |  | 15.85 | 14.1804 | 17.2186 | 3.4 | 17.4 | 14.95 | 15.5 | 14.9 | 15.45 | 16.05 | 15.6868 | 12.45 | 15.05 | 15.9 | 6
| Std Dev. |  | 0.2429 | 0.22797 | 0.56407 | 1.5002 | 0.35024 | 0.28577 | 0.49261 | 0.47188 | 0.45019 | 0.6186 | 0.22445 | 0.31623 | 1.3054 | 0.36009 | 0.63246
| Rel.Std.Dev. |  | 1.542% | 1.604% | 3.305% | 45.01% | 1.998% | 1.916% | 3.206% | 3.181% | 2.911% | 3.826% | 1.436% | 2.55% | 8.645% | 2.262% | 10.54%
| PDM3 |  | 0.3345% | -9.467% | 8.723% | -78.77% | 11.7% | -4.974% | -2.107% | -5.505% | -1.47% | 2.989% | -0.4556% | -21.01% | -3.806% | 1.396% | -61.78%
| Z-Score (Absolute) |  | 0.41 | 1.5 | 2.05 | 15.04 | 2.63 | 0.62 | 0.06 | 0.73 | 0.06 | 0.93 | 0.26 | 3.76 | 0.39 | 0.62 | NA
Results from laboratory 1.28 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C965: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1020: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 263b (Indicative Value 12.69 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.68 | 1
| 1 | 2 | 12.69 | 9
| Mean |  | 12.685 | 16
| Median |  | 12.685 | 12.685
| Std Dev. |  | 0.0070711 | 15
| Rel.Std.Dev. |  | 0.05574%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 263b (Indicative Value 275 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 270 | 1
| 1 | 2 | 280 | 10
| Mean |  | 275 | 16
| Median |  | 275 | 275
| Std Dev. |  | 7.0711 | 16
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 263b (Indicative Value 542 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 547 | 1
| 1 | 2 | 536 | 11
| Mean |  | 541.5 | 16
| Median |  | 541.5 | 541.5025
| Std Dev. |  | 7.7782 | 17
| Rel.Std.Dev. |  | 1.436%
| PDM3 |  | -0.0004617%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 263b (Indicative Value 0.41 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.41 | 1
| 1 | 2 | 0.41 | 12
| Mean |  | 0.41 | 16
| Median |  | 0.41 | 0.41
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 263b (Indicative Value 40 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 30 | 13
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 14.1421 | 19
| Rel.Std.Dev. |  | 35.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 263b (Indicative Value 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 14
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 263b (Indicative Value 146 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 146 | 1
| 1 | 2 | 146 | 15
| Mean |  | 146 | 16
| Median |  | 146 | 146.15
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.1026%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 263b (Indicative Value 85 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 90 | 16
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 263b (Indicative Value 4.02 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.003 | 1
| 1 | 2 | 4.046 | 9
| Mean |  | 4.0245 | 16
| Median |  | 4.0245 | 4.0246
| Std Dev. |  | 0.030406 | 15
| Rel.Std.Dev. |  | 0.7555%
| PDM3 |  | -0.002621%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 263b (Indicative Value 2.52 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.52 | 1
| 1 | 2 | 2.52 | 10
| Mean |  | 2.52 | 16
| Median |  | 2.52 | 2.52
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 263b (Indicative Value 0.955 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.96 | 11
| Mean |  | 0.955 | 16
| Median |  | 0.955 | 0.955
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.7404%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 263b (Indicative Value 0.025 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.025 | 1
| 1 | 2 | 0.025 | 12
| Mean |  | 0.025 | 16
| Median |  | 0.025 | 0.025
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 263b (Indicative Value 0.175 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.18 | 1
| 1 | 2 | 0.17 | 13
| Mean |  | 0.175 | 16
| Median |  | 0.175 | 0.175
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 263b (Indicative Value 70 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 70 | 14
| Mean |  | 70 | 16
| Median |  | 70 | 70
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 263b (Indicative Value 0.07 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0687 | 1
| 1 | 2 | 0.071 | 15
| Mean |  | 0.06985 | 16
| Median |  | 0.06985 | 0.069894
| Std Dev. |  | 0.0016263 | 21
| Rel.Std.Dev. |  | 2.328%
| PDM3 |  | -0.06267%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 263b (Indicative Value 50 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 16
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 263b (Indicative Value 0.027 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.027 | 1
| 1 | 2 | 0.027 | 20
| Mean |  | 0.027 | 16
| Median |  | 0.027 | 0.027
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 263b (Indicative Value 75.62 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 75.64 | 1
| 1 | 2 | 75.59 | 10
| Mean |  | 75.615 | 16
| Median |  | 75.615 | 75.615
| Std Dev. |  | 0.035355 | 16
| Rel.Std.Dev. |  | 0.04676%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 263b (Indicative Value 65 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 60 | 11
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 7.0711 | 17
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 263b (Indicative Value 50 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 40 | 1
| 1 | 2 | 60 | 12
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 14.1421 | 18
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 263b (Indicative Value 0.642 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.639 | 1
| 1 | 2 | 0.645 | 13
| Mean |  | 0.642 | 16
| Median |  | 0.642 | 0.642
| Std Dev. |  | 0.0042426 | 19
| Rel.Std.Dev. |  | 0.6608%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 263b (Indicative Value 161 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 161 | 1
| 1 | 2 | 161 | 14
| Mean |  | 161 | 16
| Median |  | 161 | 160.668
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.2066%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 263b (Indicative Value 45 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 40 | 15
| Mean |  | 45 | 16
| Median |  | 45 | 45
| Std Dev. |  | 7.0711 | 21
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 263b (Indicative Value 240 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 240 | 1
| 1 | 2 | 240 | 16
| Mean |  | 240 | 16
| Median |  | 240 | 240
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 263b (Indicative Value 2.67 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.67 | 1
| 1 | 2 | 2.67 | 18
| Mean |  | 2.67 | 16
| Median |  | 2.67 | 2.67
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 263b (Indicative Value 0.06 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.07 | 20
| Mean |  | 0.06 | 16
| Median |  | 0.06 | 0.06
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 23.57%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 263b (Indicative Value 0.155 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.16 | 20
| Mean |  | 0.155 | 16
| Median |  | 0.155 | 0.155
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 263b (Indicative Value 0.3 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.3 | 22
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 263b (Indicative Value 260 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 258 | 1
| 1 | 2 | 262 | 10
| Mean |  | 260 | 16
| Median |  | 260 | 260
| Std Dev. |  | 2.8284 | 29
| Rel.Std.Dev. |  | 1.088%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 263b (Indicative Value 494 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 494 | 1
| 1 | 2 | 494 | 24
| Mean |  | 494 | 16
| Median |  | 494 | 494
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 263b (Indicative Value 2.4 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.6 | 25
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0.28284 | 31
| Rel.Std.Dev. |  | 11.79%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 263b (Indicative Value 12 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12 | 1
| 1 | 2 | 12 | 26
| Mean |  | 12 | 16
| Median |  | 12 | 12
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 263b (Indicative Value 0.075 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.1 | 1
| 1 | 2 | 0.1 | 27
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.075
| Std Dev. |  |  | 33
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 263b (Indicative Value 84 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 83.5 | 1
| 1 | 2 | 84.6 | 28
| Mean |  | 84.05 | 16
| Median |  | 84.05 | 84.05
| Std Dev. |  | 0.77782 | 34
| Rel.Std.Dev. |  | 0.9254%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 263b (Indicative Value 5.3 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.2 | 1
| 1 | 2 | 5.4 | 14
| Mean |  | 5.3 | 16
| Median |  | 5.3 | 5.3
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 263b (Indicative Value 98 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 98 | 1
| 1 | 2 | 98 | 30
| Mean |  | 98 | 16
| Median |  | 98 | 98
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 263b (Indicative Value 8.22 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.23 | 1
| 1 | 2 | 8.2 | 31
| Mean |  | 8.215 | 16
| Median |  | 8.215 | 8.215
| Std Dev. |  | 0.021213 | 37
| Rel.Std.Dev. |  | 0.2582%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 263b (Indicative Value 74 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72 | 1
| 1 | 2 | 76 | 16
| Mean |  | 74 | 16
| Median |  | 74 | 74
| Std Dev. |  | 2.8284 | 38
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 263b (Indicative Value 4.88 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.87 | 1
| 1 | 2 | 4.88 | 4
| Mean |  | 4.875 | 16
| Median |  | 4.875 | 4.875
| Std Dev. |  | 0.0070711 | 39
| Rel.Std.Dev. |  | 0.145%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 263b (Indicative Value 2.9 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.9 | 1
| 1 | 2 | 2.89 | 5
| Mean |  | 2.895 | 16
| Median |  | 2.895 | 2.895
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.2443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 263b (Indicative Value 1.07 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.04 | 1
| 1 | 2 | 1.1 | 35
| Mean |  | 1.07 | 16
| Median |  | 1.07 | 1.07
| Std Dev. |  | 0.042426 | 41
| Rel.Std.Dev. |  | 3.965%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 263b (Indicative Value 16.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.3 | 1
| 1 | 2 | 16.3 | 36
| Mean |  | 16.3 | 16
| Median |  | 16.3 | 16.3
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 263b (Indicative Value 5.18 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.22 | 1
| 1 | 2 | 5.14 | 6
| Mean |  | 5.18 | 16
| Median |  | 5.18 | 5.18
| Std Dev. |  | 0.056569 | 43
| Rel.Std.Dev. |  | 1.092%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 263b (Indicative Value 1.6 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.55 | 1
| 1 | 2 | 1.65 | 7
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0.070711 | 44
| Rel.Std.Dev. |  | 4.419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 263b (Indicative Value 6.98 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.84 | 1
| 1 | 2 | 7.12 | 22
| Mean |  | 6.98 | 16
| Median |  | 6.98 | 6.98
| Std Dev. |  | 0.19799 | 28
| Rel.Std.Dev. |  | 2.837%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 263b (Indicative Value 1.03 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.03 | 1
| 1 | 2 | 1.02 | 23
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 263b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 4
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 263b (Indicative Value 42.4 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 42.8 | 1
| 1 | 2 | 42 | 25
| Mean |  | 42.4 | 16
| Median |  | 42.4 | 42.4
| Std Dev. |  | 0.56569 | 31
| Rel.Std.Dev. |  | 1.334%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 263b (Indicative Value 0.4 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.39 | 1
| 1 | 2 | 0.4 | 5
| Mean |  | 0.395 | 16
| Median |  | 0.395 | 0.395
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 263b (Indicative Value 0.018 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0179 | 1
| 1 | 2 | 0.0178 | 27
| Mean |  | 0.01785 | 16
| Median |  | 0.01785 | 0.01785
| Std Dev. |  | 7.0711e-05 | 33
| Rel.Std.Dev. |  | 0.3961%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 263b (Indicative Value 2.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.6 | 28
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 263b (Indicative Value 13.3 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.1 | 1
| 1 | 2 | 13.4 | 29
| Mean |  | 13.25 | 16
| Median |  | 13.25 | 13.25
| Std Dev. |  | 0.21213 | 35
| Rel.Std.Dev. |  | 1.601%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 263b (Indicative Value 34.8 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 34.7 | 1
| 1 | 2 | 34.9 | 6
| Mean |  | 34.8 | 16
| Median |  | 34.8 | 34.8
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 0.4064%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 263b (Indicative Value 46 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 46 | 14
| Mean |  | 46 | 16
| Median |  | 46 | 46
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 263b (Indicative Value 20 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 16
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 263b (Indicative Value 9.62 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.7 | 1
| 1 | 2 | 9.54 | 33
| Mean |  | 9.62 | 16
| Median |  | 9.62 | 9.62
| Std Dev. |  | 0.11314 | 39
| Rel.Std.Dev. |  | 1.176%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 263b (Indicative Value 129 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 127 | 1
| 1 | 2 | 130 | 34
| Mean |  | 128.5 | 16
| Median |  | 128.5 | 128.5
| Std Dev. |  | 2.1213 | 40
| Rel.Std.Dev. |  | 1.651%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 263b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 2
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 263b (Indicative Value 13.4 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.5 | 1
| 1 | 2 | 13.2 | 36
| Mean |  | 13.35 | 16
| Median |  | 13.35 | 13.35
| Std Dev. |  | 0.21213 | 42
| Rel.Std.Dev. |  | 1.589%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 263b (Indicative Value 11.6 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.6 | 1
| 1 | 2 | 11.5 | 37
| Mean |  | 11.55 | 16
| Median |  | 11.55 | 11.55
| Std Dev. |  | 0.070711 | 43
| Rel.Std.Dev. |  | 0.6122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 263b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 38
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 263b (Indicative Value 6.54 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.45 | 1
| 1 | 2 | 6.62 | 5
| Mean |  | 6.535 | 16
| Median |  | 6.535 | 6.535
| Std Dev. |  | 0.12021 | 28
| Rel.Std.Dev. |  | 1.839%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 263b (Indicative Value 14.6 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.8 | 1
| 1 | 2 | 14.4 | 11
| Mean |  | 14.6 | 16
| Median |  | 14.6 | 14.6
| Std Dev. |  | 0.28284 | 29
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 263b (Indicative Value 55 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 54.7 | 1
| 1 | 2 | 55.1 | 12
| Mean |  | 54.9 | 16
| Median |  | 54.9 | 54.9
| Std Dev. |  | 0.28284 | 30
| Rel.Std.Dev. |  | 0.5152%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 263b (Indicative Value 1.14 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.15 | 1
| 1 | 2 | 1.13 | 25
| Mean |  | 1.14 | 16
| Median |  | 1.14 | 1.14
| Std Dev. |  | 0.014142 | 31
| Rel.Std.Dev. |  | 1.241%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 263b (Indicative Value 0.81 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.82 | 26
| Mean |  | 0.81 | 16
| Median |  | 0.81 | 0.81
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 263b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 263b (Indicative Value 15.6 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.3 | 1
| 1 | 2 | 15.9 | 28
| Mean |  | 15.6 | 16
| Median |  | 15.6 | 15.6
| Std Dev. |  | 0.42426 | 34
| Rel.Std.Dev. |  | 2.72%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 263b (Indicative Value 0.396 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.393 | 1
| 1 | 2 | 0.399 | 29
| Mean |  | 0.396 | 16
| Median |  | 0.396 | 0.396
| Std Dev. |  | 0.0042426 | 35
| Rel.Std.Dev. |  | 1.071%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 263b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 30
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 263b (Indicative Value 0.45 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.44 | 1
| 1 | 2 | 0.45 | 6
| Mean |  | 0.445 | 16
| Median |  | 0.445 | 0.445
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 1.589%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 263b (Indicative Value 2.88 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.89 | 1
| 1 | 2 | 2.86 | 32
| Mean |  | 2.875 | 16
| Median |  | 2.875 | 2.875
| Std Dev. |  | 0.021213 | 38
| Rel.Std.Dev. |  | 0.7379%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 263b (Indicative Value 87 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 86.8 | 1
| 1 | 2 | 86.2 | 33
| Mean |  | 86.5 | 16
| Median |  | 86.5 | 86.5
| Std Dev. |  | 0.42426 | 39
| Rel.Std.Dev. |  | 0.4905%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 263b (Indicative Value 46.3 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 47 | 1
| 1 | 2 | 45.5 | 34
| Mean |  | 46.25 | 16
| Median |  | 46.25 | 46.25
| Std Dev. |  | 1.0607 | 40
| Rel.Std.Dev. |  | 2.293%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 263b (Indicative Value 26.5 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.5 | 1
| 1 | 2 | 26.4 | 35
| Mean |  | 26.45 | 16
| Median |  | 26.45 | 26.45
| Std Dev. |  | 0.070711 | 41
| Rel.Std.Dev. |  | 0.2673%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 263b (Indicative Value 2.87 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.83 | 1
| 1 | 2 | 2.9 | 7
| Mean |  | 2.865 | 16
| Median |  | 2.865 | 2.865
| Std Dev. |  | 0.049497 | 42
| Rel.Std.Dev. |  | 1.728%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 263b (Indicative Value 45 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 50 | 15
| Mean |  | 45 | 16
| Median |  | 45 | 45
| Std Dev. |  | 7.0711 | 43
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 263b (Indicative Value 243 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 241 | 1
| 1 | 2 | 245 | 16
| Mean |  | 243 | 16
| Median |  | 243 | 243
| Std Dev. |  | 2.8284 | 44
| Rel.Std.Dev. |  | 1.164%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.