## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 926b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 10.9946 | 10.8406 | 11.1486 | 10.851 | 11.1382
BaO, Barium oxide (wt.%) | 0.19311 | 0.18181 | 0.20441 | 0.1819 | 0.20431
CaO, Calcium oxide (wt.%) | 0.050124 | 0.038234 | 0.062015 | 0.047663 | 0.052585
Cu, Copper (wt.%) | 0.87812 | 0.85164 | 0.90459 | 0.86688 | 0.88935
Fe2O3, Iron(III) oxide (wt.%) | 5.3164 | 5.2385 | 5.3944 | 5.2686 | 5.3643
K2O, Potassium oxide (wt.%) | 2.9473 | 2.9074 | 2.9873 | 2.9181 | 2.9766
MgO, Magnesium oxide (wt.%) | 0.56215 | 0.53694 | 0.58736 | 0.54415 | 0.58015
P2O5, Phosphorus(V) oxide (wt.%) | 0.057097 | 0.043394 | 0.0708 | IND | IND
S, Sulphur (wt.%) | 0.97539 | 0.93619 | 1.0146 | 0.95569 | 0.99509
SiO2, Silicon dioxide (wt.%) | 74.5299 | 73.5999 | 75.4599 | 74.1548 | 74.905
TiO2, Titanium dioxide (wt.%) | 0.47667 | 0.45447 | 0.49887 | 0.46003 | 0.49331
Zn, Zinc (ppm) | 102.0629 | 81.932 | 122.1938 | IND | IND
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 3.4353 | 3.318 | 3.5526 | 3.3942 | 3.4764
Infrared Combustion
C, Carbon (wt.%) | 0.79916 | 0.7758 | 0.82252 | 0.78459 | 0.81373
S, Sulphur (wt.%) | 0.98564 | 0.95973 | 1.0115 | 0.97196 | 0.99931
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.7814 | 5.6338 | 5.9291 | 5.6622 | 5.9006
As, Arsenic (ppm) | 102.2085 | 95.0228 | 109.3943 | 96.7613 | 107.6558
B, Boron (ppm) | 130.6111 | 110.6169 | 150.6054 | 123.9432 | 137.279
Ba, Barium (wt.%) | 0.17682 | 0.16798 | 0.18565 | 0.17251 | 0.18113
Be, Beryllium (ppm) | 2.0403 | 1.95 | 2.1306 | IND | IND
Bi, Bismuth (ppm) | 49.7053 | 46.9183 | 52.4924 | 48.3781 | 51.0326
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (ppm) | 67.3343 | 63.4392 | 71.2294 | 65.4167 | 69.252
Co, Cobalt (ppm) | 7.2365 | 6.6805 | 7.7925 | 6.7963 | 7.6767
Cr, Chromium (ppm) | 83.7043 | 69.2651 | 98.1435 | 72.6777 | 94.7309
Cs, Caesium (ppm) | 4.456 | 4.1119 | 4.8002 | 4.1304 | 4.7816
Cu, Copper (wt.%) | 0.86478 | 0.84486 | 0.88469 | 0.8511 | 0.87846
Dy, Dysprosium (ppm) | 3.8955 | 3.5719 | 4.2191 | 3.756 | 4.035
Er, Erbium (ppm) | 2.3956 | 2.2269 | 2.5644 | 2.2684 | 2.5228
Eu, Europium (ppm) | 1.1025 | 0.98796 | 1.217 | 0.96508 | 1.2399
Fe, Iron (wt.%) | 3.6903 | 3.5786 | 3.8019 | 3.6269 | 3.7536
Ga, Gallium (ppm) | 16.8884 | 16.0503 | 17.7265 | 16.4118 | 17.3651
Gd, Gadolinium (ppm) | 4.6863 | 4.3369 | 5.0358 | 4.397 | 4.9756
Ge, Germanium (ppm) | 2.0751 | 1.8594 | 2.2908 | IND | IND
Ho, Holmium (ppm) | 0.80529 | 0.76046 | 0.85011 | 0.76726 | 0.84332
In, Indium (ppm) | 0.802 | 0.71036 | 0.89363 | IND | IND
K, Potassium (wt.%) | 2.4609 | 2.3623 | 2.5594 | 2.395 | 2.5267
La, Lanthanum (ppm) | 35.6649 | 33.7871 | 37.5426 | 34.7686 | 36.5612
Li, Lithium (ppm) | 24.3571 | 21.6304 | 27.0839 | 22.0328 | 26.6815
Lu, Lutetium (ppm) | 0.3452 | 0.29613 | 0.39427 | 0.32564 | 0.36476
Mg, Magnesium (wt.%) | 0.33164 | 0.31858 | 0.34469 | 0.32297 | 0.34031
Mn, Manganese (wt.%) | 0.0074128 | 0.0070267 | 0.0077989 | 0.0070017 | 0.0078239
Mo, Molybdenum (ppm) | 5.0157 | 4.8098 | 5.2217 | IND | IND
Nb, Niobium (ppm) | 10.218 | 8.5599 | 11.8761 | 9.2653 | 11.1707
Nd, Neodymium (ppm) | 29.2436 | 27.1782 | 31.3089 | 28.1124 | 30.3747
P, Phosphorus (wt.%) | 0.028574 | 0.022578 | 0.034571 | IND | IND
Pb, Lead (ppm) | 51.0648 | 44.4322 | 57.6973 | 46.6495 | 55.4801
Pr, Praseodymium (ppm) | 7.845 | 7.2949 | 8.3952 | 7.5133 | 8.1768
Rb, Rubidium (ppm) | 140.5121 | 133.1929 | 147.8313 | 137.6927 | 143.3315
S, Sulphur (wt.%) | 0.98054 | 0.94047 | 1.0206 | 0.9585 | 1.0026
Sb, Antimony (ppm) | 4.9767 | 4.1855 | 5.7679 | 4.1964 | 5.7569
Si, Silicon (wt.%) | 35.241 | 34.3323 | 36.1496 | 34.7102 | 35.7717
Sm, Samarium (ppm) | 5.5091 | 4.9807 | 6.0376 | 5.1873 | 5.831
Sn, Tin (ppm) | 13.3433 | 11.2676 | 15.419 | IND | IND
Sr, Strontium (ppm) | 45.6259 | 41.6249 | 49.627 | 43.6573 | 47.5945
Tb, Terbium (ppm) | 0.69121 | 0.6568 | 0.72561 | 0.66247 | 0.71994
Th, Thorium (ppm) | 12.6358 | 11.9106 | 13.361 | 12.2723 | 12.9992
Ti, Titanium (wt.%) | 0.28116 | 0.27174 | 0.29058 | 0.27334 | 0.28898
Tl, Thallium (ppm) | 0.97399 | 0.86895 | 1.079 | IND | IND
Tm, Thulium (ppm) | 0.35483 | 0.32091 | 0.38876 | 0.32631 | 0.38335
U, Uranium (ppm) | 4.2017 | 3.8588 | 4.5445 | 4.018 | 4.3853
V, Vanadium (ppm) | 176.5835 | 168.9818 | 184.1852 | 170.9108 | 182.2563
W, Tungsten (ppm) | 6.5039 | 4.9685 | 8.0394 | IND | IND
Y, Yttrium (ppm) | 23.1075 | 22.114 | 24.101 | 22.453 | 23.762
Yb, Ytterbium (ppm) | 2.3043 | 2.1913 | 2.4173 | 2.1486 | 2.46
Zn, Zinc (ppm) | 109.9748 | 100.3535 | 119.5961 | 101.4125 | 118.5372
4-Acid Digestion
Ag, Silver (ppm) | 3.9593 | 3.69 | 4.2286 | 3.7668 | 4.1519
Al, Aluminium (wt.%) | 5.6932 | 5.5533 | 5.8331 | 5.558 | 5.8285
As, Arsenic (ppm) | 100.2843 | 96.5035 | 104.0651 | 97.2285 | 103.3401
Ba, Barium (wt.%) | 0.16785 | 0.1575 | 0.17819 | 0.15885 | 0.17684
Be, Beryllium (ppm) | 1.9926 | 1.8477 | 2.1374 | 1.907 | 2.0781
Bi, Bismuth (ppm) | 48.0083 | 44.8536 | 51.163 | 46.7911 | 49.2255
Ca, Calcium (wt.%) | 0.037763 | 0.030694 | 0.044832 | 0.035881 | 0.039645
Cd, Cadmium (ppm) | 0.16255 | 0.1099 | 0.2152 | 0.13769 | 0.1874
Ce, Cerium (ppm) | 67.3413 | 63.7555 | 70.9271 | 65.3655 | 69.3171
Co, Cobalt (ppm) | 7.0075 | 6.4187 | 7.5962 | 6.678 | 7.337
Cr, Chromium (ppm) | 73.879 | 69.8072 | 77.9507 | 71.7181 | 76.0399
Cs, Caesium (ppm) | 4.2004 | 3.9674 | 4.4333 | 4.0447 | 4.3561
Cu, Copper (wt.%) | 0.85457 | 0.8269 | 0.88224 | 0.83936 | 0.86977
Dy, Dysprosium (ppm) | 2.9328 | 2.5919 | 3.2737 | 2.7598 | 3.1058
Er, Erbium (ppm) | 1.4932 | 1.2277 | 1.7587 | 1.3609 | 1.6255
Eu, Europium (ppm) | 1.111 | 0.97957 | 1.2424 | 1.0356 | 1.1864
Fe, Iron (wt.%) | 3.686 | 3.5668 | 3.8052 | 3.6086 | 3.7633
Ga, Gallium (ppm) | 16.53 | 15.7099 | 17.3501 | 16.0164 | 17.0435
Gd, Gadolinium (ppm) | 4.2662 | 3.914 | 4.6183 | 4.1056 | 4.4267
Hf, Hafnium (ppm) | 2.4768 | 2.2821 | 2.6715 | 2.3392 | 2.6144
Ho, Holmium (ppm) | 0.53762 | 0.44301 | 0.63223 | 0.48244 | 0.59279
In, Indium (ppm) | 0.75514 | 0.70371 | 0.80657 | 0.72167 | 0.78862
K, Potassium (wt.%) | 2.4462 | 2.3756 | 2.5169 | 2.3856 | 2.5068
La, Lanthanum (ppm) | 34.8033 | 33.2137 | 36.3929 | 33.3625 | 36.2442
Li, Lithium (ppm) | 21.507 | 20.5575 | 22.4566 | 20.7555 | 22.2586
Lu, Lutetium (ppm) | 0.21256 | 0.17833 | 0.24678 | IND | IND
Mg, Magnesium (wt.%) | 0.32634 | 0.3137 | 0.33897 | 0.31729 | 0.33538
Mn, Manganese (wt.%) | 0.0073895 | 0.0070587 | 0.0077204 | 0.0071383 | 0.0076408
Mo, Molybdenum (ppm) | 4.9405 | 4.5272 | 5.3539 | 4.6645 | 5.2166
Na, Sodium (wt.%) | 0.090217 | 0.087323 | 0.093112 | 0.087594 | 0.09284
Nb, Niobium (ppm) | 4.7212 | 3.9845 | 5.4579 | 4.4187 | 5.0238
Nd, Neodymium (ppm) | 29.1395 | 26.9305 | 31.3484 | 27.9888 | 30.2901
Ni, Nickel (ppm) | 8.8378 | 8.2524 | 9.4233 | 8.4776 | 9.1981
P, Phosphorus (wt.%) | 0.025417 | 0.024141 | 0.026693 | 0.024494 | 0.02634
Pb, Lead (ppm) | 51.1048 | 45.2158 | 56.9938 | 43.9128 | 58.2969
Pr, Praseodymium (ppm) | 7.9523 | 7.301 | 8.6036 | 7.6393 | 8.2653
Rb, Rubidium (ppm) | 142.1221 | 134.541 | 149.7032 | 137.3131 | 146.9311
S, Sulphur (wt.%) | 0.98809 | 0.96253 | 1.0136 | 0.96225 | 1.0139
Sb, Antimony (ppm) | 4.6262 | 4.2777 | 4.9748 | 4.3763 | 4.8761
Sc, Scandium (ppm) | 9.692 | 9.1178 | 10.2661 | 9.2932 | 10.0907
Se, Selenium (ppm) | 11.7677 | 10.697 | 12.8385 | 11.049 | 12.4864
Sm, Samarium (ppm) | 5.5611 | 5.1169 | 6.0053 | 5.3018 | 5.8203
Sn, Tin (ppm) | 10.4575 | 9.7745 | 11.1406 | 9.9559 | 10.9591
Sr, Strontium (ppm) | 40.2433 | 39.0345 | 41.4521 | 38.9379 | 41.5486
Tb, Terbium (ppm) | 0.55619 | 0.47823 | 0.63416 | 0.50554 | 0.60684
Te, Tellurium (ppm) | 0.077259 | 0.051404 | 0.10311 | 0.062876 | 0.091642
Th, Thorium (ppm) | 12.3984 | 11.7908 | 13.006 | 11.9697 | 12.827
Ti, Titanium (wt.%) | 0.20256 | 0.18487 | 0.22026 | 0.19639 | 0.20874
Tl, Thallium (ppm) | 0.9405 | 0.88001 | 1.001 | 0.87196 | 1.009
U, Uranium (ppm) | 3.9733 | 3.752 | 4.1947 | 3.8107 | 4.136
V, Vanadium (ppm) | 168.955 | 163.757 | 174.153 | 163.8803 | 174.0297
W, Tungsten (ppm) | 6.0067 | 5.4693 | 6.5441 | 5.7525 | 6.2609
Y, Yttrium (ppm) | 12.8213 | 11.6028 | 14.0399 | 12.1816 | 13.4611
Yb, Ytterbium (ppm) | 1.5038 | 1.2376 | 1.77 | 1.3698 | 1.6378
Zn, Zinc (ppm) | 104.1496 | 93.9857 | 114.3135 | 94.9861 | 113.3132
Zr, Zirconium (ppm) | 83.1132 | 78.8514 | 87.3749 | 79.9723 | 86.254
Aqua Regia Digestion
Ag, Silver (ppm) | 3.7834 | 3.5687 | 3.9982 | 3.6565 | 3.9103
Al, Aluminium (wt.%) | 0.47935 | 0.4304 | 0.52831 | 0.46474 | 0.49396
As, Arsenic (ppm) | 96.4435 | 92.5103 | 100.3766 | 94.5477 | 98.3393
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.31271 | 0.27078 | 0.35463 | 0.29435 | 0.33107
Bi, Bismuth (ppm) | 47.8303 | 45.7228 | 49.9377 | 46.1188 | 49.5417
Ca, Calcium (wt.%) | 0.030293 | 0.02669 | 0.033897 | IND | IND
Cd, Cadmium (ppm) | 0.15479 | 0.11892 | 0.19065 | 0.14128 | 0.16829
Ce, Cerium (ppm) | 49.2938 | 44.4079 | 54.1796 | 47.5436 | 51.0439
Co, Cobalt (ppm) | 6.7782 | 6.4146 | 7.1418 | 6.5658 | 6.9905
Cr, Chromium (ppm) | 21.8831 | 20.7974 | 22.9688 | 20.9716 | 22.7945
Cs, Caesium (ppm) | 1.0095 | 0.88545 | 1.1335 | 0.96255 | 1.0564
Cu, Copper (wt.%) | 0.87334 | 0.84943 | 0.89725 | 0.86076 | 0.88592
Dy, Dysprosium (ppm) | 1.5232 | 1.1957 | 1.8508 | 1.3985 | 1.648
Fe, Iron (wt.%) | 3.4751 | 3.3948 | 3.5555 | 3.4082 | 3.5421
Ga, Gallium (ppm) | 1.6734 | 1.4604 | 1.8864 | 1.5373 | 1.8095
Gd, Gadolinium (ppm) | 2.8116 | 2.0287 | 3.5945 | 2.603 | 3.0202
Ge, Germanium (ppm) | 0.09959 | 0.072076 | 0.1271 | IND | IND
Hf, Hafnium (ppm) | 0.28674 | 0.26538 | 0.30811 | 0.26714 | 0.30634
Hg, Mercury (ppm) | 0.038648 | 0.025156 | 0.05214 | IND | IND
In, Indium (ppm) | 0.72106 | 0.66247 | 0.77965 | 0.69956 | 0.74256
K, Potassium (wt.%) | 0.17412 | 0.15896 | 0.18928 | 0.16743 | 0.18081
La, Lanthanum (ppm) | 22.4408 | 19.6129 | 25.2687 | 21.4785 | 23.4032
Li, Lithium (ppm) | 1.6108 | 1.2967 | 1.9249 | IND | IND
Mg, Magnesium (wt.%) | 0.046893 | 0.039719 | 0.054068 | IND | IND
Mn, Manganese (wt.%) | 0.005155 | 0.0049593 | 0.0053506 | 0.0049981 | 0.0053118
Mo, Molybdenum (ppm) | 4.7689 | 4.4874 | 5.0504 | 4.5907 | 4.9471
Ni, Nickel (ppm) | 6.8989 | 6.3646 | 7.4333 | 6.4499 | 7.348
P, Phosphorus (wt.%) | 0.020526 | 0.019355 | 0.021698 | 0.019941 | 0.021112
Pb, Lead (ppm) | 40.3209 | 35.6132 | 45.0287 | 35.1398 | 45.5021
Rb, Rubidium (ppm) | 10.1878 | 8.9746 | 11.401 | 9.6798 | 10.6958
S, Sulphur (wt.%) | 0.99324 | 0.95554 | 1.0309 | 0.97167 | 1.0148
Sb, Antimony (ppm) | 3.0309 | 2.6952 | 3.3666 | 2.8196 | 3.2422
Sc, Scandium (ppm) | 1.2727 | 1.1418 | 1.4035 | IND | IND
Se, Selenium (ppm) | 11.3764 | 10.2395 | 12.5132 | 10.8692 | 11.8836
Sm, Samarium (ppm) | 3.6185 | 2.6696 | 4.5675 | 3.3149 | 3.9222
Sn, Tin (ppm) | 7.5017 | 7.187 | 7.8163 | 7.258 | 7.7454
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Te, Tellurium (ppm) | 0.069985 | 0.051367 | 0.088603 | IND | IND
Th, Thorium (ppm) | 8.0296 | 7.4975 | 8.5617 | 7.7033 | 8.3559
Tl, Thallium (ppm) | 0.093617 | 0.075989 | 0.11124 | IND | IND
U, Uranium (ppm) | 1.9227 | 1.8107 | 2.0348 | 1.8578 | 1.9877
V, Vanadium (ppm) | 14.0718 | 12.2847 | 15.8589 | 13.6295 | 14.514
W, Tungsten (ppm) | 3.1781 | 2.8261 | 3.5301 | 3.0259 | 3.3303
Y, Yttrium (ppm) | 5.0439 | 4.5272 | 5.5606 | 4.8298 | 5.258
Zn, Zinc (ppm) | 96.8167 | 86.0925 | 107.5408 | 86.5634 | 107.07
Zr, Zirconium (ppm) | 8.7863 | 8.0783 | 9.4944 | 8.3614 | 9.2113
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 926b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
As | ppm | 141.6667 | Mo | ppm | < 50 | Sm | ppm | < 90
Bi | ppm | 96.6667 | Na2O | wt.% | 0.18366 | Sn | ppm | < 50
CeO2 | ppm | < 100 | Nb | ppm | < 50 | Sr | ppm | 100
Co | ppm | < 100 | Nd | ppm | < 90 | Ta | ppm | 200
Cr | ppm | 72.369 | Ni | ppm | < 50 | Th | ppm | < 44
Gd | ppm | < 90 | Pb | ppm | 108.1481 | V | ppm | 197.3336
HfO2 | ppm | < 100 | Pr | ppm | 74.491 | W | ppm | 15.4167
La2O3 | ppm | 75 | Rb | ppm | 151.2761 | Y2O3 | ppm | 66.6667
MnO2 | wt.% | 0.015549 | Sb | ppm | 192 | Zr | ppm | 144.4444
Borate / Peroxide Fusion ICP
Ag | ppm | 3.4075 | Ni | ppm | 16.9911 | Ta | ppm | 1.0556
Ca | wt.% | 0.067418 | Re | ppm | < 0.1 | Te | ppm | < 5
Hf | ppm | 4.1571 | Sc | ppm | 9.0298 | Zr | ppm | 146.8583
Na | wt.% | 0.084075 | Se | ppm | 10.3004
4-Acid Digestion
B | ppm | 25.1667 | Hg | ppm | 0.032 | Ta | ppm | 0.32188
Ge | ppm | 0.16256 | Re | ppb | 1.7185 | Tm | ppm | 0.1975
Aqua Regia Digestion
Au | ppm | 0.017214 | Na | wt.% | 0.010765 | Re | ppb | 0.88333
Ba | ppm | 105.3058 | Nb | ppm | 0.0442 | Sr | ppm | 8.1509
Er | ppm | 0.54 | Nd | ppm | 20.4295 | Tb | ppm | 0.29017
Eu | ppm | 0.60458 | Pd | ppb | 31.25 | Ti | ppm | 48.8549
Ho | ppm | 0.23312 | Pr | ppm | 5.5529 | Tm | ppm | 0.069444
Lu | ppm | 0.0675 | Pt | ppb | < 10 | Yb | ppm | 0.48042
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 926b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 10.9946 | 0.17188 | 10.6509 | 11.3384 | 10.479 | 11.5102 | 1.563% | 3.127% | 4.69% | 10.4449 | 11.5443
BaO, wt.% | 0.19311 | 0.010776 | 0.17156 | 0.21466 | 0.16078 | 0.22544 | 5.58% | 11.16% | 16.74% | 0.18345 | 0.20276
CaO, wt.% | 0.050124 | 0.011976 | 0.026173 | 0.074076 | 0.014198 | 0.086051 | 23.89% | 47.78% | 71.68% | 0.047618 | 0.052631
Cu, wt.% | 0.87812 | 0.037022 | 0.80407 | 0.95216 | 0.76705 | 0.98918 | 4.216% | 8.432% | 12.65% | 0.83421 | 0.92202
Fe2O3, wt.% | 5.3164 | 0.069464 | 5.1775 | 5.4554 | 5.108 | 5.5248 | 1.307% | 2.613% | 3.92% | 5.0506 | 5.5822
K2O, wt.% | 2.9473 | 0.044885 | 2.8576 | 3.0371 | 2.8127 | 3.082 | 1.523% | 3.046% | 4.569% | 2.8 | 3.0947
MgO, wt.% | 0.56215 | 0.031565 | 0.49902 | 0.62528 | 0.46746 | 0.65685 | 5.615% | 11.23% | 16.84% | 0.53404 | 0.59026
P2O5, wt.% | 0.057097 | 0.010529 | 0.036039 | 0.078155 | 0.02551 | 0.088683 | 18.44% | 36.88% | 55.32% | 0.054242 | 0.059952
S, wt.% | 0.97539 | 0.052794 | 0.8698 | 1.081 | 0.81701 | 1.1338 | 5.413% | 10.83% | 16.24% | 0.92662 | 1.0242
SiO2, wt.% | 74.5299 | 0.7174 | 73.0951 | 75.9647 | 72.3777 | 76.6821 | 0.9626% | 1.925% | 2.888% | 70.8034 | 78.2564
TiO2, wt.% | 0.47667 | 0.0088409 | 0.45898 | 0.49435 | 0.45014 | 0.50319 | 1.855% | 3.709% | 5.564% | 0.45283 | 0.5005
Zn, ppm | 102.0629 | 13.6561 | 74.7508 | 129.3751 | 61.0947 | 143.0312 | 13.38% | 26.76% | 40.14% | 96.9598 | 107.1661
Thermogravimetry
LOI1000, wt.% | 3.4353 | 0.13255 | 3.1702 | 3.7004 | 3.0376 | 3.8329 | 3.859% | 7.717% | 11.58% | 3.2635 | 3.607
Infrared Combustion
C, wt.% | 0.79916 | 0.037146 | 0.72487 | 0.87345 | 0.68772 | 0.9106 | 4.648% | 9.296% | 13.94% | 0.7592 | 0.83912
S, wt.% | 0.98564 | 0.032022 | 0.92159 | 1.0497 | 0.88957 | 1.0817 | 3.249% | 6.498% | 9.746% | 0.93636 | 1.0349
Borate / Peroxide Fusion ICP
Al, wt.% | 5.7814 | 0.093749 | 5.5939 | 5.9689 | 5.5002 | 6.0627 | 1.622% | 3.243% | 4.865% | 5.4924 | 6.0705
As, ppm | 102.2085 | 3.9503 | 94.3079 | 110.1092 | 90.3575 | 114.0596 | 3.865% | 7.73% | 11.59% | 97.0981 | 107.319
B, ppm | 130.6111 | 15.6271 | 99.357 | 161.8652 | 83.7299 | 177.4923 | 11.96% | 23.93% | 35.89% | 124.0806 | 137.1417
Ba, wt.% | 0.17682 | 0.011873 | 0.15307 | 0.20056 | 0.1412 | 0.21244 | 6.715% | 13.43% | 20.14% | 0.16798 | 0.18566
Be, ppm | 2.0403 | 0.072409 | 1.8955 | 2.1851 | 1.8231 | 2.2575 | 3.549% | 7.098% | 10.65% | 1.9383 | 2.1423
Bi, ppm | 49.7053 | 2.2973 | 45.1107 | 54.2999 | 42.8134 | 56.5972 | 4.622% | 9.244% | 13.87% | 47.2201 | 52.1906
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 67.3343 | 4.7153 | 57.9038 | 76.7648 | 53.1886 | 81.4801 | 7.003% | 14.01% | 21.01% | 63.9676 | 70.701
Co, ppm | 7.2365 | 0.32722 | 6.5821 | 7.8909 | 6.2549 | 8.2182 | 4.522% | 9.044% | 13.57% | 6.8747 | 7.5983
Cr, ppm | 83.7043 | 12.5148 | 58.6747 | 108.734 | 46.1599 | 121.2488 | 14.95% | 29.9% | 44.85% | 79.5191 | 87.8895
Cs, ppm | 4.456 | 0.22261 | 4.0108 | 4.9012 | 3.7882 | 5.1239 | 4.996% | 9.991% | 14.99% | 4.2332 | 4.6788
Cu, wt.% | 0.86478 | 0.017941 | 0.8289 | 0.90066 | 0.81096 | 0.9186 | 2.075% | 4.149% | 6.224% | 0.82154 | 0.90802
Dy, ppm | 3.8955 | 0.37004 | 3.1554 | 4.6356 | 2.7854 | 5.0056 | 9.499% | 19% | 28.5% | 3.7007 | 4.0903
Er, ppm | 2.3956 | 0.12028 | 2.1551 | 2.6362 | 2.0348 | 2.7565 | 5.021% | 10.04% | 15.06% | 2.2758 | 2.5154
Eu, ppm | 1.1025 | 0.051686 | 0.99913 | 1.2059 | 0.94744 | 1.2576 | 4.688% | 9.376% | 14.06% | 1.0474 | 1.1576
Fe, wt.% | 3.6903 | 0.10275 | 3.4848 | 3.8958 | 3.382 | 3.9985 | 2.784% | 5.569% | 8.353% | 3.5057 | 3.8748
Ga, ppm | 16.8884 | 0.47386 | 15.9407 | 17.8361 | 15.4668 | 18.31 | 2.806% | 5.612% | 8.418% | 16.044 | 17.7328
Gd, ppm | 4.6863 | 0.26469 | 4.157 | 5.2157 | 3.8923 | 5.4804 | 5.648% | 11.3% | 16.94% | 4.452 | 4.9206
Ge, ppm | 2.0751 | 0.11438 | 1.8464 | 2.3039 | 1.732 | 2.4182 | 5.512% | 11.02% | 16.54% | 1.9714 | 2.1789
Ho, ppm | 0.80529 | 0.029969 | 0.74535 | 0.86522 | 0.71538 | 0.89519 | 3.722% | 7.443% | 11.16% | 0.76502 | 0.84555
In, ppm | 0.802 | 0.044067 | 0.71386 | 0.89013 | 0.6698 | 0.9342 | 5.495% | 10.99% | 16.48% | 0.7619 | 0.8421
K, wt.% | 2.4609 | 0.081918 | 2.297 | 2.6247 | 2.2151 | 2.7066 | 3.329% | 6.658% | 9.986% | 2.3378 | 2.5839
La, ppm | 35.6649 | 1.7502 | 32.1646 | 39.1652 | 30.4144 | 40.9153 | 4.907% | 9.814% | 14.72% | 33.8816 | 37.4481
Li, ppm | 24.3571 | 4.3119 | 15.7334 | 32.9809 | 11.4216 | 37.2927 | 17.7% | 35.41% | 53.11% | 23.1393 | 25.575
Lu, ppm | 0.3452 | 0.025104 | 0.29499 | 0.39541 | 0.26989 | 0.42051 | 7.272% | 14.54% | 21.82% | 0.32794 | 0.36246
Mg, wt.% | 0.33164 | 0.013806 | 0.30402 | 0.35925 | 0.29022 | 0.37305 | 4.163% | 8.326% | 12.49% | 0.31505 | 0.34822
Mn, wt.% | 0.0074128 | 0.00053167 | 0.0063495 | 0.0084761 | 0.0058178 | 0.0090078 | 7.172% | 14.34% | 21.52% | 0.0070422 | 0.0077834
Mo, ppm | 5.0157 | 0.34172 | 4.3323 | 5.6991 | 3.9905 | 6.0409 | 6.813% | 13.63% | 20.44% | 4.7649 | 5.2665
Nb, ppm | 10.218 | 1.3306 | 7.5568 | 12.8792 | 6.2262 | 14.2098 | 13.02% | 26.04% | 39.07% | 9.7071 | 10.7289
Nd, ppm | 29.2436 | 1.7634 | 25.7167 | 32.7704 | 23.9533 | 34.5339 | 6.03% | 12.06% | 18.09% | 27.7814 | 30.7057
P, wt.% | 0.028574 | 0.0037968 | 0.020981 | 0.036168 | 0.017184 | 0.039965 | 13.29% | 26.57% | 39.86% | 0.027146 | 0.030003
Pb, ppm | 51.0648 | 5.536 | 39.9928 | 62.1368 | 34.4568 | 67.6728 | 10.84% | 21.68% | 32.52% | 48.5115 | 53.618
Pr, ppm | 7.845 | 0.63504 | 6.575 | 9.1151 | 5.9399 | 9.7502 | 8.095% | 16.19% | 24.28% | 7.4528 | 8.2373
Rb, ppm | 140.5121 | 5.3812 | 129.7497 | 151.2744 | 124.3685 | 156.6556 | 3.83% | 7.659% | 11.49% | 133.4865 | 147.5377
S, wt.% | 0.98054 | 0.053031 | 0.87448 | 1.0866 | 0.82145 | 1.1396 | 5.408% | 10.82% | 16.22% | 0.93152 | 1.0296
Sb, ppm | 4.9767 | 0.56891 | 3.8389 | 6.1145 | 3.2699 | 6.6834 | 11.43% | 22.86% | 34.29% | 4.7278 | 5.2255
Si, wt.% | 35.241 | 0.92377 | 33.3934 | 37.0885 | 32.4696 | 38.0123 | 2.621% | 5.243% | 7.864% | 33.4789 | 37.003
Sm, ppm | 5.5091 | 0.46934 | 4.5704 | 6.4478 | 4.1011 | 6.9171 | 8.519% | 17.04% | 25.56% | 5.2337 | 5.7846
Sn, ppm | 13.3433 | 1.5314 | 10.2805 | 16.4062 | 8.749 | 17.9376 | 11.48% | 22.95% | 34.43% | 12.6762 | 14.0105
Sr, ppm | 45.6259 | 5.5055 | 34.615 | 56.6368 | 29.1096 | 62.1423 | 12.07% | 24.13% | 36.2% | 43.3446 | 47.9072
Tb, ppm | 0.69121 | 0.042798 | 0.60561 | 0.7768 | 0.56282 | 0.8196 | 6.192% | 12.38% | 18.58% | 0.65665 | 0.72577
Th, ppm | 12.6358 | 0.82279 | 10.9902 | 14.2814 | 10.1674 | 15.1042 | 6.512% | 13.02% | 19.53% | 12.004 | 13.2676
Ti, wt.% | 0.28116 | 0.0093548 | 0.26245 | 0.29987 | 0.2531 | 0.30922 | 3.327% | 6.654% | 9.982% | 0.2671 | 0.29522
Tl, ppm | 0.97399 | 0.047943 | 0.87811 | 1.0699 | 0.83016 | 1.1178 | 4.922% | 9.845% | 14.77% | 0.92529 | 1.0227
Tm, ppm | 0.35483 | 0.022966 | 0.3089 | 0.40076 | 0.28594 | 0.42373 | 6.472% | 12.94% | 19.42% | 0.33709 | 0.37258
U, ppm | 4.2017 | 0.18401 | 3.8336 | 4.5697 | 3.6496 | 4.7537 | 4.38% | 8.759% | 13.14% | 3.9916 | 4.4118
V, ppm | 176.5835 | 7.61 | 161.3636 | 191.8034 | 153.7536 | 199.4134 | 4.31% | 8.619% | 12.93% | 167.7543 | 185.4127
W, ppm | 6.5039 | 1.1381 | 4.2277 | 8.7801 | 3.0896 | 9.9182 | 17.5% | 35% | 52.5% | 6.1787 | 6.8291
Y, ppm | 23.1075 | 0.63993 | 21.8276 | 24.3873 | 21.1877 | 25.0273 | 2.769% | 5.539% | 8.308% | 21.9521 | 24.2628
Yb, ppm | 2.3043 | 0.14434 | 2.0157 | 2.593 | 1.8713 | 2.7373 | 6.264% | 12.53% | 18.79% | 2.1891 | 2.4195
Zn, ppm | 109.9748 | 9.7729 | 90.429 | 129.5206 | 80.6562 | 139.2935 | 8.886% | 17.77% | 26.66% | 104.4761 | 115.4735
4-Acid Digestion
Ag, ppm | 3.9593 | 0.28266 | 3.394 | 4.5246 | 3.1113 | 4.8073 | 7.139% | 14.28% | 21.42% | 3.7613 | 4.1573
Al, wt.% | 5.6932 | 0.16448 | 5.3643 | 6.0222 | 5.1998 | 6.1867 | 2.889% | 5.778% | 8.667% | 5.4086 | 5.9779
As, ppm | 100.2843 | 3.8165 | 92.6512 | 107.9173 | 88.8347 | 111.7339 | 3.806% | 7.611% | 11.42% | 95.2701 | 105.2985
Ba, wt.% | 0.16785 | 0.01217 | 0.14351 | 0.19219 | 0.13134 | 0.20436 | 7.251% | 14.5% | 21.75% | 0.15945 | 0.17624
Be, ppm | 1.9926 | 0.11058 | 1.7714 | 2.2137 | 1.6608 | 2.3243 | 5.55% | 11.1% | 16.65% | 1.8929 | 2.0922
Bi, ppm | 48.0083 | 3.5006 | 41.007 | 55.0096 | 37.5064 | 58.5102 | 7.292% | 14.58% | 21.88% | 45.6079 | 50.4087
Ca, wt.% | 0.037763 | 0.0060362 | 0.02569 | 0.049835 | 0.019654 | 0.055872 | 15.98% | 31.97% | 47.95% | 0.035875 | 0.039651
Cd, ppm | 0.16255 | 0.039459 | 0.083628 | 0.24146 | 0.044169 | 0.28092 | 24.28% | 48.55% | 72.83% | 0.15442 | 0.17067
Ce, ppm | 67.3413 | 4.556 | 58.2293 | 76.4534 | 53.6733 | 81.0094 | 6.766% | 13.53% | 20.3% | 63.9743 | 70.7084
Co, ppm | 7.0075 | 0.51321 | 5.981 | 8.0339 | 5.4678 | 8.5471 | 7.324% | 14.65% | 21.97% | 6.6571 | 7.3578
Cr, ppm | 73.879 | 5.6455 | 62.588 | 85.1699 | 56.9426 | 90.8154 | 7.642% | 15.28% | 22.92% | 70.185 | 77.5729
Cs, ppm | 4.2004 | 0.15838 | 3.8836 | 4.5171 | 3.7253 | 4.6755 | 3.77% | 7.541% | 11.31% | 3.9904 | 4.4104
Cu, wt.% | 0.85457 | 0.026257 | 0.80205 | 0.90708 | 0.7758 | 0.93334 | 3.073% | 6.145% | 9.218% | 0.81184 | 0.8973
Dy, ppm | 2.9328 | 0.37497 | 2.1829 | 3.6827 | 1.8079 | 4.0577 | 12.79% | 25.57% | 38.36% | 2.7862 | 3.0795
Er, ppm | 1.4932 | 0.25862 | 0.97593 | 2.0104 | 0.71731 | 2.269 | 17.32% | 34.64% | 51.96% | 1.4185 | 1.5678
Eu, ppm | 1.111 | 0.049085 | 1.0128 | 1.2092 | 0.96375 | 1.2583 | 4.418% | 8.836% | 13.25% | 1.0555 | 1.1666
Fe, wt.% | 3.686 | 0.14484 | 3.3963 | 3.9756 | 3.2514 | 4.1205 | 3.929% | 7.859% | 11.79% | 3.5017 | 3.8703
Ga, ppm | 16.53 | 0.74848 | 15.033 | 18.0269 | 14.2845 | 18.7754 | 4.528% | 9.056% | 13.58% | 15.7035 | 17.3565
Gd, ppm | 4.2662 | 0.28195 | 3.7023 | 4.8301 | 3.4203 | 5.112 | 6.609% | 13.22% | 19.83% | 4.0529 | 4.4795
Hf, ppm | 2.4768 | 0.16008 | 2.1566 | 2.797 | 1.9966 | 2.957 | 6.463% | 12.93% | 19.39% | 2.353 | 2.6006
Ho, ppm | 0.53762 | 0.094363 | 0.34889 | 0.72635 | 0.25453 | 0.82071 | 17.55% | 35.1% | 52.66% | 0.51074 | 0.5645
In, ppm | 0.75514 | 0.039904 | 0.67533 | 0.83495 | 0.63543 | 0.87486 | 5.284% | 10.57% | 15.85% | 0.71739 | 0.7929
K, wt.% | 2.4462 | 0.067026 | 2.3122 | 2.5803 | 2.2452 | 2.6473 | 2.74% | 5.48% | 8.22% | 2.3239 | 2.5685
La, ppm | 34.8033 | 1.6713 | 31.4607 | 38.1459 | 29.7894 | 39.8172 | 4.802% | 9.604% | 14.41% | 33.0632 | 36.5435
Li, ppm | 21.507 | 0.8425 | 19.822 | 23.192 | 18.9795 | 24.0345 | 3.917% | 7.835% | 11.75% | 20.4317 | 22.5824
Lu, ppm | 0.21256 | 0.015107 | 0.18234 | 0.24277 | 0.16723 | 0.25788 | 7.108% | 14.22% | 21.32% | 0.20193 | 0.22318
Mg, wt.% | 0.32634 | 0.015301 | 0.29573 | 0.35694 | 0.28043 | 0.37224 | 4.689% | 9.378% | 14.07% | 0.31002 | 0.34265
Mn, wt.% | 0.0073895 | 0.00039875 | 0.006592 | 0.008187 | 0.0061933 | 0.0085858 | 5.396% | 10.79% | 16.19% | 0.0070201 | 0.007759
Mo, ppm | 4.9405 | 0.297 | 4.3465 | 5.5345 | 4.0495 | 5.8315 | 6.011% | 12.02% | 18.03% | 4.6935 | 5.1876
Na, wt.% | 0.090217 | 0.0028371 | 0.084543 | 0.095891 | 0.081706 | 0.098729 | 3.145% | 6.29% | 9.434% | 0.085706 | 0.094728
Nb, ppm | 4.7212 | 0.92078 | 2.8797 | 6.5628 | 1.9589 | 7.4836 | 19.5% | 39.01% | 58.51% | 4.4852 | 4.9573
Nd, ppm | 29.1395 | 1.5351 | 26.0694 | 32.2096 | 24.5343 | 33.7446 | 5.268% | 10.54% | 15.8% | 27.6825 | 30.5964
Ni, ppm | 8.8378 | 0.79849 | 7.2408 | 10.4348 | 6.4423 | 11.2333 | 9.035% | 18.07% | 27.1% | 8.3959 | 9.2797
P, wt.% | 0.025417 | 0.0011317 | 0.023154 | 0.027681 | 0.022022 | 0.028812 | 4.453% | 8.905% | 13.36% | 0.024146 | 0.026688
Pb, ppm | 51.1048 | 7.7476 | 35.6096 | 66.6001 | 27.862 | 74.3477 | 15.16% | 30.32% | 45.48% | 48.5496 | 53.6601
Pr, ppm | 7.9523 | 0.44856 | 7.0552 | 8.8494 | 6.6066 | 9.298 | 5.641% | 11.28% | 16.92% | 7.5547 | 8.3499
Rb, ppm | 142.1221 | 6.3634 | 129.3953 | 154.8489 | 123.0319 | 161.2123 | 4.477% | 8.955% | 13.43% | 135.016 | 149.2282
S, wt.% | 0.98809 | 0.036807 | 0.91448 | 1.0617 | 0.87767 | 1.0985 | 3.725% | 7.45% | 11.18% | 0.93868 | 1.0375
Sb, ppm | 4.6262 | 0.35991 | 3.9064 | 5.346 | 3.5465 | 5.7059 | 7.78% | 15.56% | 23.34% | 4.3949 | 4.8575
Sc, ppm | 9.692 | 0.74725 | 8.1975 | 11.1865 | 7.4502 | 11.9337 | 7.71% | 15.42% | 23.13% | 9.2074 | 10.1766
Se, ppm | 11.7677 | 0.88505 | 9.9976 | 13.5378 | 9.1126 | 14.4229 | 7.521% | 15.04% | 22.56% | 11.1793 | 12.3561
Sm, ppm | 5.5611 | 0.34637 | 4.8683 | 6.2538 | 4.522 | 6.6002 | 6.228% | 12.46% | 18.69% | 5.283 | 5.8391
Sn, ppm | 10.4575 | 0.41957 | 9.6184 | 11.2967 | 9.1988 | 11.7162 | 4.012% | 8.024% | 12.04% | 9.9346 | 10.9804
Sr, ppm | 40.2433 | 1.3025 | 37.6382 | 42.8483 | 36.3357 | 44.1508 | 3.237% | 6.473% | 9.71% | 38.2311 | 42.2554
Tb, ppm | 0.55619 | 0.053857 | 0.44848 | 0.66391 | 0.39462 | 0.71776 | 9.683% | 19.37% | 29.05% | 0.52838 | 0.584
Te, ppm | 0.077259 | 0.015133 | 0.046993 | 0.10753 | 0.03186 | 0.12266 | 19.59% | 39.17% | 58.76% | 0.073396 | 0.081122
Th, ppm | 12.3984 | 0.85407 | 10.6902 | 14.1065 | 9.8362 | 14.9606 | 6.889% | 13.78% | 20.67% | 11.7784 | 13.0183
Ti, wt.% | 0.20256 | 0.029253 | 0.14406 | 0.26107 | 0.1148 | 0.29032 | 14.44% | 28.88% | 43.32% | 0.19243 | 0.21269
Tl, ppm | 0.9405 | 0.066657 | 0.80719 | 1.0738 | 0.74053 | 1.1405 | 7.087% | 14.17% | 21.26% | 0.89347 | 0.98752
U, ppm | 3.9733 | 0.23066 | 3.512 | 4.4347 | 3.2814 | 4.6653 | 5.805% | 11.61% | 17.42% | 3.7747 | 4.172
V, ppm | 168.955 | 5.6919 | 157.5712 | 180.3388 | 151.8793 | 186.0307 | 3.369% | 6.738% | 10.11% | 160.5073 | 177.4028
W, ppm | 6.0067 | 0.54744 | 4.9118 | 7.1016 | 4.3644 | 7.649 | 9.114% | 18.23% | 27.34% | 5.7064 | 6.307
Y, ppm | 12.8213 | 1.4234 | 9.9745 | 15.6682 | 8.5511 | 17.0916 | 11.1% | 22.2% | 33.31% | 12.1803 | 13.4624
Yb, ppm | 1.5038 | 0.26958 | 0.96464 | 2.0429 | 0.69506 | 2.3125 | 17.93% | 35.85% | 53.78% | 1.4286 | 1.579
Zn, ppm | 104.1496 | 13.1878 | 77.774 | 130.5253 | 64.5861 | 143.7131 | 12.66% | 25.32% | 37.99% | 98.9422 | 109.3571
Zr, ppm | 83.1132 | 5.3379 | 72.4373 | 93.7891 | 67.0993 | 99.127 | 6.423% | 12.85% | 19.27% | 78.9575 | 87.2688
Aqua Regia Digestion
Ag, ppm | 3.7834 | 0.1808 | 3.4218 | 4.145 | 3.241 | 4.3258 | 4.779% | 9.558% | 14.34% | 3.5942 | 3.9726
Al, wt.% | 0.47935 | 0.091408 | 0.29654 | 0.66217 | 0.20513 | 0.75358 | 19.07% | 38.14% | 57.21% | 0.45539 | 0.50332
As, ppm | 96.4435 | 3.1213 | 90.201 | 102.686 | 87.0797 | 105.8073 | 3.236% | 6.473% | 9.709% | 91.6213 | 101.2657
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.31271 | 0.04537 | 0.22197 | 0.40345 | 0.1766 | 0.44882 | 14.51% | 29.02% | 43.53% | 0.29707 | 0.32834
Bi, ppm | 47.8303 | 2.2757 | 43.2789 | 52.3816 | 41.0032 | 54.6573 | 4.758% | 9.516% | 14.27% | 45.4387 | 50.2218
Ca, wt.% | 0.030293 | 0.003545 | 0.023203 | 0.037383 | 0.019658 | 0.040928 | 11.7% | 23.4% | 35.11% | 0.028779 | 0.031808
Cd, ppm | 0.15479 | 0.025738 | 0.10331 | 0.20626 | 0.077575 | 0.232 | 16.63% | 33.26% | 49.88% | 0.14705 | 0.16253
Ce, ppm | 49.2938 | 8.4271 | 32.4396 | 66.1479 | 24.0126 | 74.575 | 17.1% | 34.19% | 51.29% | 46.8291 | 51.7585
Co, ppm | 6.7782 | 0.30203 | 6.1741 | 7.3822 | 5.8721 | 7.6843 | 4.456% | 8.912% | 13.37% | 6.4393 | 7.1171
Cr, ppm | 21.8831 | 1.5356 | 18.8119 | 24.9543 | 17.2763 | 26.4899 | 7.017% | 14.03% | 21.05% | 20.7889 | 22.9772
Cs, ppm | 1.0095 | 0.12453 | 0.76039 | 1.2585 | 0.63586 | 1.3831 | 12.34% | 24.67% | 37.01% | 0.95898 | 1.0599
Cu, wt.% | 0.87334 | 0.026159 | 0.82102 | 0.92566 | 0.79486 | 0.95182 | 2.995% | 5.991% | 8.986% | 0.82967 | 0.91701
Dy, ppm | 1.5232 | 0.21058 | 1.1021 | 1.9444 | 0.89151 | 2.155 | 13.82% | 27.65% | 41.47% | 1.4471 | 1.5994
Fe, wt.% | 3.4751 | 0.075805 | 3.3235 | 3.6268 | 3.2477 | 3.7026 | 2.181% | 4.363% | 6.544% | 3.3014 | 3.6489
Ga, ppm | 1.6734 | 0.28101 | 1.1114 | 2.2354 | 0.83041 | 2.5165 | 16.79% | 33.58% | 50.38% | 1.5898 | 1.7571
Gd, ppm | 2.8116 | 0.56177 | 1.6881 | 3.9351 | 1.1263 | 4.4969 | 19.98% | 39.96% | 59.94% | 2.671 | 2.9522
Ge, ppm | 0.09959 | 0.026533 | 0.046525 | 0.15265 | 0.019992 | 0.17919 | 26.64% | 53.28% | 79.93% | 0.09461 | 0.10457
Hf, ppm | 0.28674 | 0.017399 | 0.25194 | 0.32154 | 0.23454 | 0.33894 | 6.068% | 12.14% | 18.2% | 0.2724 | 0.30108
Hg, ppm | 0.038648 | 0.0080456 | 0.022557 | 0.054739 | 0.014511 | 0.062785 | 20.82% | 41.63% | 62.45% | 0.036716 | 0.040581
In, ppm | 0.72106 | 0.039393 | 0.64228 | 0.79985 | 0.60288 | 0.83924 | 5.463% | 10.93% | 16.39% | 0.68501 | 0.75711
K, wt.% | 0.17412 | 0.02099 | 0.13214 | 0.2161 | 0.11115 | 0.23709 | 12.05% | 24.11% | 36.16% | 0.16541 | 0.18282
La, ppm | 22.4408 | 4.8664 | 12.7081 | 32.1735 | 7.8417 | 37.0399 | 21.69% | 43.37% | 65.06% | 21.3188 | 23.5629
Li, ppm | 1.6108 | 0.38999 | 0.83083 | 2.3908 | 0.44084 | 2.7808 | 24.21% | 48.42% | 72.63% | 1.5303 | 1.6914
Mg, wt.% | 0.046893 | 0.007536 | 0.031821 | 0.061965 | 0.024285 | 0.069501 | 16.07% | 32.14% | 48.21% | 0.044549 | 0.049238
Mn, wt.% | 0.005155 | 0.00019887 | 0.0047572 | 0.0055527 | 0.0045584 | 0.0057516 | 3.858% | 7.716% | 11.57% | 0.0048972 | 0.0054127
Mo, ppm | 4.7689 | 0.29912 | 4.1706 | 5.3671 | 3.8715 | 5.6662 | 6.272% | 12.54% | 18.82% | 4.5304 | 5.0073
Ni, ppm | 6.8989 | 0.35922 | 6.1805 | 7.6174 | 5.8213 | 7.9766 | 5.207% | 10.41% | 15.62% | 6.554 | 7.2439
P, wt.% | 0.020526 | 0.00094417 | 0.018638 | 0.022415 | 0.017694 | 0.023359 | 4.6% | 9.2% | 13.8% | 0.0195 | 0.021553
Pb, ppm | 40.3209 | 5.5895 | 29.142 | 51.4999 | 23.5525 | 57.0894 | 13.86% | 27.72% | 41.59% | 38.3049 | 42.337
Rb, ppm | 10.1878 | 1.5536 | 7.0805 | 13.295 | 5.5269 | 14.8486 | 15.25% | 30.5% | 45.75% | 9.6784 | 10.6972
S, wt.% | 0.99324 | 0.041388 | 0.91047 | 1.076 | 0.86908 | 1.1174 | 4.167% | 8.334% | 12.5% | 0.94358 | 1.0429
Sb, ppm | 3.0309 | 0.58708 | 1.8567 | 4.2051 | 1.2697 | 4.7921 | 19.37% | 38.74% | 58.11% | 2.8794 | 3.1825
Sc, ppm | 1.2727 | 0.13429 | 1.0041 | 1.5412 | 0.8698 | 1.6755 | 10.55% | 21.1% | 31.65% | 1.209 | 1.3363
Se, ppm | 11.3764 | 1.1102 | 9.1559 | 13.5969 | 8.0456 | 14.7071 | 9.759% | 19.52% | 29.28% | 10.8075 | 11.9452
Sm, ppm | 3.6185 | 0.70329 | 2.212 | 5.0251 | 1.5087 | 5.7284 | 19.44% | 38.87% | 58.31% | 3.4376 | 3.7995
Sn, ppm | 7.5017 | 0.26806 | 6.9656 | 8.0378 | 6.6975 | 8.3058 | 3.573% | 7.147% | 10.72% | 7.1266 | 7.8767
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Te, ppm | 0.069985 | 0.011373 | 0.047239 | 0.092731 | 0.035866 | 0.1041 | 16.25% | 32.5% | 48.75% | 0.066486 | 0.073484
Th, ppm | 8.0296 | 0.86565 | 6.2983 | 9.7609 | 5.4326 | 10.6265 | 10.78% | 21.56% | 32.34% | 7.6281 | 8.4311
Tl, ppm | 0.093617 | 0.0093477 | 0.074921 | 0.11231 | 0.065574 | 0.12166 | 9.985% | 19.97% | 29.96% | 0.088936 | 0.098297
U, ppm | 1.9227 | 0.1377 | 1.6473 | 2.1981 | 1.5096 | 2.3358 | 7.162% | 14.32% | 21.48% | 1.8266 | 2.0189
V, ppm | 14.0718 | 2.9787 | 8.1144 | 20.0292 | 5.1357 | 23.0079 | 21.17% | 42.34% | 63.5% | 13.3682 | 14.7754
W, ppm | 3.1781 | 0.51097 | 2.1562 | 4.2001 | 1.6452 | 4.711 | 16.08% | 32.16% | 48.23% | 3.0192 | 3.337
Y, ppm | 5.0439 | 0.80412 | 3.4357 | 6.6521 | 2.6315 | 7.4562 | 15.94% | 31.88% | 47.83% | 4.7917 | 5.2961
Zn, ppm | 96.8167 | 16.6257 | 63.5653 | 130.068 | 46.9396 | 146.6937 | 17.17% | 34.34% | 51.52% | 91.9758 | 101.6575
Zr, ppm | 8.7863 | 0.73179 | 7.3228 | 10.2499 | 6.591 | 10.9817 | 8.329% | 16.66% | 24.99% | 8.347 | 9.2256
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 926b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 926b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Inspectorate Griffith India Pvt. Ltd., Zawar, Rajasthan, India
Intertek, Perth, WA, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion, South Africa

## Fusion XRF
| Analytical results for Al2O3 in OREAS 926b (Certified Value 10.99 wt.%) | [18 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 10.934 | 11 | 10.9 | 11.15 | 11.15 | 11.2 | 12.155 | 11.09 | 10.58 | 10.8445 | 11.26 | 10.8 | 10.9 | 7.9188 | 10.83 | 11.0357 |  | 1
| 1 | 2 | 11.013 | 11.1 | 10.9 | 11.2 | 11.1 | 11.2 | 12.195 | 11.11 | 10.7 | 10.8251 | 11.18 | 10.8 | 11.3 | 8.1904 | 10.81 | 11.0516 |  | #N/A
| 1 | 3 | 10.989 | 11.05 | 11 | 11.35 | 11.1 | 11.1 | 10.959 | 11.09 | 10.71 | 10.8351 | 11.04 | 10.8 | 11.1 | 7.9597 | 10.69 | 11.0754 |  | 16
| 1 | 4 | 10.831 | 11.15 | 10.9 | 11.2 | 11.05 | 11.15 | 10.776 | 11.07 | 10.74 | 10.9334 | 11.18 | 10.9 | 11 | 8.0542 | 10.69 | 11.0879 |  | 10.9946
| 1 | 5 | 10.831 | 11.1 | 10.95 | 11.4 | 11 | 11.05 | 10.926 | 11.09 | 10.65 | 10.839 | 11.16 | 10.8 | 11.2 | 7.919 | 10.82 | 11.0757 |  | 7
| 1 | 6 | 11.056 | 11.15 | 11.05 | 11.3 | 11.1 | 11.15 | 10.871 | 11.08 | 10.87 | 10.9476 | 11.07 | 10.7 | 11.1 | 8.1205 | 10.81 | 11.092
| Mean |  | 10.9423 | 11.0917 | 10.95 | 11.2667 | 11.0833 | 11.1417 | 11.3137 | 11.0883 | 10.7083 | 10.8708 | 11.1483 | 10.8 | 11.1 | 8.0271 | 10.775 | 11.0697
| Median |  | 10.9615 | 11.1 | 10.925 | 11.25 | 11.1 | 11.15 | 10.9425 | 11.09 | 10.705 | 10.8417 | 11.17 | 10.8 | 11.1 | 8.0069 | 10.81 | 11.0755
| Std Dev. |  | 0.094817 | 0.058452 | 0.063246 | 0.098319 | 0.05164 | 0.058452 | 0.67018 | 0.013292 | 0.09704 | 0.054558 | 0.080602 | 0.063246 | 0.14142 | 0.11321 | 0.066257 | 0.021806
| Rel.Std.Dev. |  | 0.8665% | 0.527% | 0.5776% | 0.8727% | 0.4659% | 0.5246% | 5.924% | 0.1199% | 0.9062% | 0.5019% | 0.723% | 0.5856% | 1.274% | 1.41% | 0.6149% | 0.197%
| PDM3 |  | -0.4755% | 0.8828% | -0.4057% | 2.474% | 0.807% | 1.338% | 2.902% | 0.8524% | -2.604% | -1.126% | 1.398% | -1.77% | 0.9586% | -26.99% | -1.997% | 0.683%
| Z-Score (Absolute) |  | 0.69 | 0.08 | 0.65 | 0.98 | 0.04 | 0.34 | 1.23 | 0.06 | 1.9 | 1.06 | 0.37 | 1.43 | 0.12 | 15.77 | 1.56 | 0.04
| Analytical results for As in OREAS 926b (Indicative Value 142 ppm) | [18 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | 200 | < 100 | 200 | [14 empty] | 1
| 1 | 2 | 100 | < 100 | 200 | [14 empty] | 1
| 1 | 3 | < 100 | < 100 | 100 | [14 empty] | 16
| 1 | 4 | < 100 | < 100 | 200 | [14 empty] | 141.6667
| 1 | 5 | 100 | < 100 | 200 | [14 empty] | 7
| 1 | 6 | 100 | < 100 | 200
| Mean |  | 125 |  | 183.3333
| Median |  | 100 |  | 200
| Std Dev. |  | 50 |  | 40.8248
| Rel.Std.Dev. |  | 40% |  | 22.27%
| PDM3 |  | -11.76% |  | 29.41%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.12
| Analytical results for BaO in OREAS 926b (Certified Value 0.193 wt.%) | [18 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 19 | 20 | 21 |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.19996 | 0.2 | 0.201 | 0.19 | 0.2 | 0.19 | 0.185 | 0.2 | 0.17 | 0.19 | 0.19 | 0.2 | 0.19059 | 0.18 | 0.18175 |  |  | 1
| 1 | 2 | 0.1975 | 0.21 | 0.201 | 0.2 | 0.2 | 0.18 | 0.187 | 0.21 | 0.17 | 0.2 | 0.19 | 0.19 | 0.21638 | 0.18 | 0.19202 |  |  | #N/A
| 1 | 3 | 0.19996 | 0.2 | 0.201 | 0.2 | 0.2 | 0.18 | 0.19 | 0.2 | 0.19 | 0.18 | 0.19 | 0.19 | 0.22244 | 0.18 | 0.19758 |  |  | 16
| 1 | 4 | 0.19132 | 0.2 | 0.201 | 0.2 | 0.2 | 0.18 | 0.186 | 0.2 | 0.18 | 0.19 | 0.179 | 0.2 | 0.21832 | 0.18 | 0.18749 |  |  | 0.19311
| 1 | 5 | 0.19009 | 0.21 | 0.19 | 0.2 | 0.21 | 0.18 | 0.186 | 0.2 | 0.17 | 0.2 | 0.19 | 0.2 | 0.11688 | 0.18 | 0.18343 |  |  | 8
| 1 | 6 | 0.19996 | 0.2 | 0.19 | 0.2 | 0.2 | 0.19 | 0.181 | 0.2 | 0.18 | 0.18 | 0.179 | 0.2 | 0.25356 | 0.18 | 0.18205
| Mean |  | 0.19647 | 0.20333 | 0.19733 | 0.19833 | 0.20167 | 0.18333 | 0.18583 | 0.20167 | 0.17667 | 0.19 | 0.18633 | 0.19667 | 0.20303 | 0.18 | 0.18739
| Median |  | 0.19873 | 0.2 | 0.201 | 0.2 | 0.2 | 0.18 | 0.186 | 0.2 | 0.175 | 0.19 | 0.19 | 0.2 | 0.21735 | 0.18 | 0.18546
| Std Dev. |  | 0.0045799 | 0.005164 | 0.0056804 | 0.0040825 | 0.0040825 | 0.005164 | 0.0029269 | 0.0040825 | 0.008165 | 0.0089443 | 0.0056804 | 0.005164 | 0.04673 | 3.0405e-17 | 0.0063449
| Rel.Std.Dev. |  | 2.331% | 2.54% | 2.879% | 2.058% | 2.024% | 2.817% | 1.575% | 2.024% | 4.622% | 4.708% | 3.049% | 2.626% | 23.02% | 1.689e-14% | 3.386%
| PDM3 |  | 1.74% | 5.295% | 2.188% | 2.706% | 4.432% | -5.062% | -3.767% | 4.432% | -8.514% | -1.609% | -3.508% | 1.843% | 5.137% | -6.788% | -2.963%
| Z-Score (Absolute) |  | 0 | 0.71 | 0.08 | 0.19 | 0.53 | 1.35 | 1.09 | 0.53 | 2.04 | 0.66 | 1.07 | 0.02 | 0.67 | 1.69 | 0.93
| Analytical results for Bi in OREAS 926b (Indicative Value 97 ppm) | [18 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | 200 | 100 | 100 | [14 empty] | 1
| 1 | 2 | 100 | 100 | 100 | [14 empty] | 2
| 1 | 3 | < 100 | 100 | 100 | [14 empty] | 16
| 1 | 4 | 100 | 100 | 100 | [14 empty] | 96.6667
| 1 | 5 | 100 | 100 | 100 | [14 empty] | 8
| 1 | 6 | 100 | 100 | 100
| Mean |  | 120 | 100 | 100
| Median |  | 100 | 100 | 100
| Std Dev. |  | 44.7214 | 0 | 0
| Rel.Std.Dev. |  | 37.27% | 0% | 0%
| PDM3 |  | 24.14% | 3.448% | 3.448%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for CaO in OREAS 926b (Certified Value 0.05 wt.%) | [18 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.050611 | 0.03 | 0.04 | 0.04 | 0.04 | 0.05 | 0.071 | 0.05 | 0.07 | 0.0507 | 0.07 | 0.02 | 0.04 | 0.1145 | 0.04 | < 0.019 |  | 1
| 1 | 2 | 0.049434 | 0.03 | 0.04 | 0.05 | 0.05 | 0.05 | 0.07 | 0.05 | 0.06 | 0.049922 | 0.07 | 0.01 | 0.04 | 0.11636 | 0.04 | < 0.019 |  | #N/A
| 1 | 3 | 0.049434 | 0.03 | 0.04 | 0.05 | 0.04 | 0.05 | 0.063 | 0.05 | 0.06 | 0.050108 | 0.07 | 0.02 | 0.03 | 0.12167 | 0.04 | < 0.019 |  | 16
| 1 | 4 | 0.049434 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.061 | 0.05 | 0.06 | 0.050641 | 0.07 | 0.01 | 0.06 | 0.11174 | 0.04 | < 0.019 |  | 0.050124
| 1 | 5 | 0.050611 | 0.04 | 0.11 | 0.05 | 0.04 | 0.05 | 0.077 | 0.05 | 0.08 | 0.050156 | 0.07 | 0.01 | 0.03 | 0.10838 | 0.04 | < 0.019 |  | 9
| 1 | 6 | 0.050611 | 0.04 | 0.12 | 0.05 | 0.04 | 0.05 | 0.06 | 0.05 | 0.07 | 0.051041 | 0.07 | 0.01 | 0.03 | 0.14193 | 0.04 | < 0.019
| Mean |  | 0.050023 | 0.035 | 0.066667 | 0.048333 | 0.043333 | 0.05 | 0.067 | 0.05 | 0.066667 | 0.050428 | 0.07 | 0.013333 | 0.038333 | 0.1191 | 0.04
| Median |  | 0.050022 | 0.035 | 0.045 | 0.05 | 0.04 | 0.05 | 0.0665 | 0.05 | 0.065 | 0.050398 | 0.07 | 0.01 | 0.035 | 0.11543 | 0.04
| Std Dev. |  | 0.00064467 | 0.0054772 | 0.037771 | 0.0040825 | 0.005164 | 7.6012e-18 | 0.0067231 | 7.6012e-18 | 0.008165 | 0.00043085 | 0 | 0.005164 | 0.01169 | 0.012046 | 0
| Rel.Std.Dev. |  | 1.289% | 15.65% | 56.66% | 8.447% | 11.92% | 1.52e-14% | 10.03% | 1.52e-14% | 12.25% | 0.8544% | 0% | 38.73% | 30.5% | 10.11% | 0%
| PDM3 |  | -0.2033% | -30.17% | 33% | -3.573% | -13.55% | -0.2482% | 33.67% | -0.2482% | 33% | 0.6059% | 39.65% | -73.4% | -23.52% | 137.6% | -20.2%
| Z-Score (Absolute) |  | 0 | 0.76 | 0.84 | 0.08 | 0.34 | 0 | 0.86 | 0 | 0.84 | 0.02 | 1.01 | 1.85 | 0.59 | 3.49 | 0.51 | 2.05
| Analytical results for CeO2 in OREAS 926b (Indicative Value < 100 ppm) | [18 empty] | Indicative
CeO2 | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 06 | 12 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 300 | [14 empty] | 1
| 1 | 2 | < 100 | < 100 | 100 | [14 empty] | 3
| 1 | 3 | 100 | < 100 | < 100 | [14 empty] | 16
| 1 | 4 | < 100 | 100 | < 100 | [14 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | [14 empty] | 9
| 1 | 6 | < 100 | < 100 | < 100
| Mean |  | 100 | 100 | 200
| Median |  | 100 | 100 | 200
| Std Dev. |  |  |  | 141.4214
| Rel.Std.Dev. |  |  |  | 70.71%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Co in OREAS 926b (Indicative Value < 100 ppm) | [18 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 05 | 06 | 18 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 50 | [13 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 50 | [13 empty] | 4
| 1 | 3 | < 100 | 100 | < 100 | < 50 | [13 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | 70 | [13 empty] | < 100
| 1 | 5 | < 100 | 100 | < 100 | 70 | [13 empty] | 10
| 1 | 6 | < 100 | 100 | < 100 | < 50
| Mean |  |  | 100 |  | 70
| Median |  |  | 100 |  | 70
| Std Dev. |  |  | 0 |  | 0
| Rel.Std.Dev. |  |  | 0% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 3.37 | 0 | 1.35
| Analytical results for Cr in OREAS 926b (Indicative Value 72 ppm) | [18 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 06 | 12 | 16 | 18 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | 100 | 100 | 100 | < 34 | < 70 | 70 | [11 empty] | 1
| 1 | 2 | 100 | 100 | 100 | 48 | < 70 | < 70 | [11 empty] | 5
| 1 | 3 | 100 | 100 | 100 | 68 | 70 | < 70 | [11 empty] | 16
| 1 | 4 | 100 | 100 | 100 | 62 | < 70 | 70 | [11 empty] | 72.369
| 1 | 5 | 100 | 100 | 100 | < 34 | < 70 | 70 | [11 empty] | 11
| 1 | 6 | 100 | 100 | < 100 | 62 | < 70 | 70
| Mean |  | 100 | 100 | 100 | 60 | 70 | 70
| Median |  | 100 | 100 | 100 | 62 | 70 | 70
| Std Dev. |  | 0 | 0 | 0 | 8.4853 |  | 0
| Rel.Std.Dev. |  | 0% | 0% | 0% | 14.14% |  | 0%
| PDM3 |  | 38.18% | 38.18% | 38.18% | -17.09% | -3.274% | -3.274%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.46 | 0.76 | 0.9 | 0.46
| Analytical results for Cu in OREAS 926b (Certified Value 0.878 wt.%) | [18 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.88374 | 0.87 | 0.89 | 0.841 | 0.87 | 0.875 | 0.915 | 0.86 | 0.8859 | 0.84393 | 0.91 | 0.9 | 0.89 | 0.93635 | 0.815 | 0.8402 |  | 1
| 1 | 2 | 0.8853 | 0.89 | 0.89 | 0.832 | 0.858 | 0.881 | 0.919 | 0.87 | 0.8819 | 0.84747 | 0.92 | 0.91 | 0.91 | 0.9634 | 0.807 | 0.82628 |  | #N/A
| 1 | 3 | 0.88764 | 0.88 | 0.91 | 0.837 | 0.842 | 0.882 | 0.911 | 0.87 | 0.8875 | 0.84636 | 0.92 | 0.91 | 0.9 | 0.94764 | 0.783 | 0.84358 |  | 16
| 1 | 4 | 0.88296 | 0.88 | 0.88 | 0.832 | 0.859 | 0.879 | 0.896 | 0.86 | 0.9474 | 0.84436 | 0.9 | 0.88 | 0.9 | 0.94147 | 0.791 | 0.83162 |  | 0.87812
| 1 | 5 | 0.88374 | 0.86 | 0.9 | 0.842 | 0.883 | 0.88 | 0.911 | 0.87 | 0.8668 | 0.84507 | 0.92 | 0.88 | 0.93 | 0.92492 | 0.799 | 0.84093 |  | 10
| 1 | 6 | 0.87906 | 0.87 | 0.89 | 0.832 | 0.86 | 0.875 | 0.924 | 0.87 | 0.8915 | 0.84808 | 0.89 | 0.89 | 0.91 | 0.9877 | 0.831 | 0.83523
| Mean |  | 0.88374 | 0.875 | 0.89333 | 0.836 | 0.862 | 0.87867 | 0.91267 | 0.86667 | 0.8935 | 0.84588 | 0.91 | 0.895 | 0.90667 | 0.95025 | 0.80433 | 0.83631
| Median |  | 0.88374 | 0.875 | 0.89 | 0.8345 | 0.8595 | 0.8795 | 0.913 | 0.87 | 0.8867 | 0.84571 | 0.915 | 0.895 | 0.905 | 0.94456 | 0.803 | 0.83771
| Std Dev. |  | 0.0028339 | 0.010488 | 0.010328 | 0.0046904 | 0.013667 | 0.0030111 | 0.0095638 | 0.005164 | 0.02775 | 0.0016951 | 0.012649 | 0.013784 | 0.013663 | 0.022334 | 0.017282 | 0.0065283
| Rel.Std.Dev. |  | 0.3207% | 1.199% | 1.156% | 0.5611% | 1.586% | 0.3427% | 1.048% | 0.5958% | 3.106% | 0.2004% | 1.39% | 1.54% | 1.507% | 2.35% | 2.149% | 0.7806%
| PDM3 |  | 0.6404% | -0.3549% | 1.733% | -4.796% | -1.835% | 0.06268% | 3.935% | -1.304% | 1.752% | -3.671% | 3.631% | 1.923% | 3.251% | 8.214% | -8.402% | -4.761%
| Z-Score (Absolute) |  | 0.08 | 0.19 | 0.37 | 1.36 | 0.58 | 0.08 | 0.95 | 0.44 | 0.37 | 1.07 | 0.87 | 0.42 | 0.77 | 2.08 | 2.33 | 1.36
| Analytical results for Fe2O3 in OREAS 926b (Certified Value 5.32 wt.%) | [18 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 5.343 | 5.22 | 5.218 | 5.276 | 5.218 | 5.29 | 5.373 | 5.29 | 5.32 | 5.2997 | 5.38 | 5.34 | 5.4 | 5.3252 | 5.3 | 5.1454 |  | 1
| 1 | 2 | 5.348 | 5.29 | 5.276 | 5.276 | 5.233 | 5.33 | 5.352 | 5.261 | 5.36 | 5.2912 | 5.36 | 5.32 | 5.37 | 5.5179 | 5.28 | 5.1448 |  | #N/A
| 1 | 3 | 5.337 | 5.25 | 5.318 | 5.333 | 5.218 | 5.27 | 5.339 | 5.261 | 5.33 | 5.3411 | 5.29 | 5.32 | 5.41 | 5.4114 | 5.28 | 5.152 |  | 16
| 1 | 4 | 5.233 | 5.29 | 5.204 | 5.29 | 5.247 | 5.29 | 5.315 | 5.247 | 5.41 | 5.2959 | 5.31 | 5.34 | 5.44 | 5.4986 | 5.25 | 5.1279 |  | 5.3164
| 1 | 5 | 5.242 | 5.26 | 5.19 | 5.333 | 5.233 | 5.29 | 5.393 | 5.261 | 5.64 | 5.3573 | 5.36 | 5.35 | 5.51 | 5.2542 | 5.31 | 5.1404 |  | 11
| 1 | 6 | 5.246 | 5.3 | 5.261 | 5.347 | 5.247 | 5.32 | 5.364 | 5.261 | 5.39 | 5.3505 | 5.28 | 5.29 | 5.47 | 5.5085 | 5.31 | 5.1277
| Mean |  | 5.2915 | 5.2683 | 5.2445 | 5.3092 | 5.2327 | 5.2983 | 5.356 | 5.2635 | 5.4083 | 5.3226 | 5.33 | 5.3267 | 5.4333 | 5.4193 | 5.2883 | 5.1397
| Median |  | 5.2915 | 5.275 | 5.2395 | 5.3115 | 5.233 | 5.29 | 5.358 | 5.261 | 5.375 | 5.3204 | 5.335 | 5.33 | 5.425 | 5.455 | 5.29 | 5.1426
| Std Dev. |  | 0.056316 | 0.030605 | 0.048948 | 0.032046 | 0.012972 | 0.022286 | 0.027232 | 0.014139 | 0.11856 | 0.030153 | 0.041952 | 0.021602 | 0.050859 | 0.10967 | 0.023166 | 0.0099338
| Rel.Std.Dev. |  | 1.064% | 0.5809% | 0.9333% | 0.6036% | 0.2479% | 0.4206% | 0.5084% | 0.2686% | 2.192% | 0.5665% | 0.7871% | 0.4056% | 0.9361% | 2.024% | 0.4381% | 0.1933%
| PDM3 |  | -0.4688% | -0.9046% | -1.353% | -0.1365% | -1.575% | -0.3403% | 0.7444% | -0.9955% | 1.729% | 0.1164% | 0.2554% | 0.1927% | 2.199% | 1.935% | -0.5284% | -3.324%
| Z-Score (Absolute) |  | 0.22 | 0.63 | 1.06 | 0.1 | 1.27 | 0.1 | 0.94 | 0.72 | 1.88 | 0.34 | 0.47 | 0.41 | 2.32 | 2.07 | 0.27 | 2.94
| Analytical results for Gd in OREAS 926b (Indicative Value < 90 ppm) | [18 empty] | Indicative
Gd | Round | Replicate | Lab | [16 empty] | 1
| No | No | 12 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 0
0
| 1 | 1 | < 90 | [16 empty] | 1
| 1 | 2 | 170 | [16 empty] | 6
| 1 | 3 | < 90 | [16 empty] | 16
| 1 | 4 | < 90 | [16 empty] | < 90
| 1 | 5 | < 90 | [16 empty] | 12
| 1 | 6 | < 90
| Mean |  | 170
| Median |  | 170
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 926b (Indicative Value < 100 ppm) | [18 empty] | Indicative
HfO2 | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 05 | 06 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | [15 empty] | 0
0
| 1 | 1 | < 100 | < 100 | [15 empty] | 1
| 1 | 2 | < 100 | < 100 | [15 empty] | 7
| 1 | 3 | < 100 | < 100 | [15 empty] | 16
| 1 | 4 | < 100 | < 100 | [15 empty] | < 100
| 1 | 5 | < 100 | < 100 | [15 empty] | 13
| 1 | 6 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K2O in OREAS 926b (Certified Value 2.95 wt.%) | [18 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 2.999 | 3 | 2.95 | 2.9 | 3 | 2.98 | 2.933 | 2.98 | 2.889 | 2.95 | 3.03 | 2.9 | 2.95 | 2.4377 | 2.95 | 2.9055 |  | 1
| 1 | 2 | 3.003 | 2.97 | 2.95 | 2.9 | 3 | 3 | 2.898 | 2.96 | 2.926 | 2.94 | 3 | 2.86 | 3.02 | 2.4997 | 2.95 | 2.9024 |  | #N/A
| 1 | 3 | 2.998 | 2.95 | 2.99 | 2.91 | 2.99 | 2.99 | 2.932 | 2.98 | 2.887 | 2.96 | 2.99 | 2.86 | 2.98 | 2.5132 | 2.94 | 2.9353 |  | 16
| 1 | 4 | 2.898 | 2.96 | 2.94 | 2.87 | 2.98 | 3.01 | 2.92 | 2.96 | 2.918 | 2.95 | 3.02 | 2.87 | 2.93 | 2.5083 | 2.94 | 2.9063 |  | 2.9473
| 1 | 5 | 2.903 | 2.87 | 2.95 | 2.91 | 2.99 | 2.99 | 2.927 | 2.98 | 2.986 | 2.94 | 3.01 | 2.87 | 2.95 | 2.4058 | 2.94 | 2.8958 |  | 12
| 1 | 6 | 2.897 | 2.98 | 2.99 | 2.9 | 3.01 | 3 | 2.864 | 2.99 | 2.926 | 2.96 | 2.95 | 2.87 | 3.02 | 2.4991 | 2.93 | 2.9011
| Mean |  | 2.9497 | 2.955 | 2.9617 | 2.8983 | 2.995 | 2.995 | 2.9123 | 2.975 | 2.922 | 2.95 | 3 | 2.8717 | 2.975 | 2.4773 | 2.9417 | 2.9077
| Median |  | 2.9505 | 2.965 | 2.95 | 2.9 | 2.995 | 2.995 | 2.9235 | 2.98 | 2.922 | 2.95 | 3.005 | 2.87 | 2.965 | 2.4994 | 2.94 | 2.904
| Std Dev. |  | 0.0552 | 0.045056 | 0.022286 | 0.01472 | 0.010488 | 0.010488 | 0.026942 | 0.012247 | 0.035939 | 0.0089443 | 0.028284 | 0.01472 | 0.038341 | 0.044513 | 0.0075277 | 0.01402
| Rel.Std.Dev. |  | 1.871% | 1.525% | 0.7525% | 0.5079% | 0.3502% | 0.3502% | 0.9251% | 0.4117% | 1.23% | 0.3032% | 0.9428% | 0.5126% | 1.289% | 1.797% | 0.2559% | 0.4822%
| PDM3 |  | 0.07903% | 0.26% | 0.4862% | -1.663% | 1.617% | 1.617% | -1.188% | 0.9386% | -0.8597% | 0.09034% | 1.787% | -2.567% | 0.9386% | -15.95% | -0.1924% | -1.344%
| Z-Score (Absolute) |  | 0 | 0.11 | 0.24 | 1.06 | 0.93 | 0.93 | 0.77 | 0.52 | 0.57 | 0 | 1.04 | 1.61 | 0.52 | 9.75 | 0.17 | 0.87
| Analytical results for La2O3 in OREAS 926b (Indicative Value 75 ppm) | [18 empty] | Indicative
La2O3 | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 06 | 12 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 200 | [14 empty] | 1
| 1 | 2 | < 100 | 100 | < 200 | [14 empty] | 8
| 1 | 3 | < 100 | 100 | < 200 | [14 empty] | 16
| 1 | 4 | < 100 | 100 | < 200 | [14 empty] | 75
| 1 | 5 | 100 | 100 | < 200 | [14 empty] | 14
| 1 | 6 | < 100 | 100 | < 200
| Mean |  | 100 | 100
| Median |  | 100 | 100
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  | 33.33% | 33.33%
| Z-Score (Absolute) |  | 2.7 | 0 | 0.67
| Analytical results for MgO in OREAS 926b (Certified Value 0.562 wt.%) | [18 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.5328 | 0.57 | 0.54 | 0.59 | 0.61 | 0.59 | 0.543 | 0.57 | 0.49 | 0.59182 | 0.58 | 0.57 | 0.49 | 0.42466 | 0.58 | 0.49072 |  | 1
| 1 | 2 | 0.5364 | 0.57 | 0.53 | 0.61 | 0.6 | 0.6 | 0.55 | 0.57 | 0.5 | 0.58478 | 0.59 | 0.56 | 0.47 | 0.41841 | 0.6 | 0.55329 |  | #N/A
| 1 | 3 | 0.5292 | 0.55 | 0.55 | 0.6 | 0.58 | 0.58 | 0.547 | 0.55 | 0.48 | 0.58108 | 0.57 | 0.57 | 0.5 | 0.42084 | 0.59 | 0.53087 |  | 16
| 1 | 4 | 0.53318 | 0.58 | 0.54 | 0.57 | 0.59 | 0.58 | 0.562 | 0.57 | 0.47 | 0.59078 | 0.58 | 0.58 | 0.54 | 0.43694 | 0.57 | 0.49629 |  | 0.56215
| 1 | 5 | 0.53212 | 0.57 | 0.56 | 0.59 | 0.59 | 0.58 | 0.56 | 0.56 | 0.5 | 0.59222 | 0.58 | 0.58 | 0.47 | 0.42707 | 0.58 | 0.49076 |  | 13
| 1 | 6 | 0.53742 | 0.58 | 0.56 | 0.58 | 0.61 | 0.58 | 0.56 | 0.58 | 0.5 | 0.58259 | 0.57 | 0.55 | 0.51 | 0.45584 | 0.59 | 0.45163
| Mean |  | 0.53352 | 0.57 | 0.54667 | 0.59 | 0.59667 | 0.585 | 0.55367 | 0.56667 | 0.49 | 0.58721 | 0.57833 | 0.56833 | 0.49667 | 0.43063 | 0.585 | 0.50226
| Median |  | 0.53299 | 0.57 | 0.545 | 0.59 | 0.595 | 0.58 | 0.555 | 0.57 | 0.495 | 0.58778 | 0.58 | 0.57 | 0.495 | 0.42586 | 0.585 | 0.49352
| Std Dev. |  | 0.002992 | 0.010954 | 0.012111 | 0.014142 | 0.012111 | 0.0083666 | 0.0080166 | 0.010328 | 0.012649 | 0.0049798 | 0.0075277 | 0.01169 | 0.026583 | 0.01392 | 0.010488 | 0.03546
| Rel.Std.Dev. |  | 0.5608% | 1.922% | 2.215% | 2.397% | 2.03% | 1.43% | 1.448% | 1.823% | 2.581% | 0.848% | 1.302% | 2.057% | 5.352% | 3.233% | 1.793% | 7.06%
| PDM3 |  | -5.093% | 1.396% | -2.755% | 4.954% | 6.14% | 4.065% | -1.509% | 0.8032% | -12.83% | 4.458% | 2.879% | 1.1% | -11.65% | -23.4% | 4.065% | -10.65%
| Z-Score (Absolute) |  | 1.13 | 0.08 | 0.69 | 0.75 | 0.97 | 0.58 | 0.46 | 0.03 | 2.58 | 0.66 | 0.36 | 0.03 | 2.36 | 4.55 | 0.58 | 2.17
| Analytical results for MnO2 in OREAS 926b (Indicative Value 0.016 wt.%) | [18 empty] | Indicative
MnO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 20 | 21 |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.012255 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0257 | < 0.012 | < 0.0061 | < 0.012 | 0.012 | 0.025 | < 0.012 | < 0.012 | < 0.0111 |  |  | 1
| 1 | 2 | 0.013481 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0245 | < 0.012 | 0.0061 | < 0.012 | < 0.012 | 0.025 | < 0.012 | 0.012 | < 0.0111 |  |  | 9
| 1 | 3 | 0.013481 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0233 | < 0.012 | < 0.0061 | < 0.012 | 0.012 | 0.025 | < 0.012 | < 0.012 | < 0.0111 |  |  | 16
| 1 | 4 | 0.014707 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0221 | < 0.012 | < 0.0061 | < 0.012 | 0.012 | 0.025 | < 0.012 | < 0.012 | < 0.0111 |  |  | 0.015549
| 1 | 5 | 0.015932 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0245 | < 0.012 | < 0.0061 | < 0.012 | 0.012 | 0.025 | < 0.012 | 0.025 | < 0.0111 |  |  | 15
| 1 | 6 | 0.013481 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0245 | < 0.012 | < 0.0061 | < 0.012 | 0.012 | 0.025 | < 0.012 | 0.025 | < 0.0111
| Mean |  | 0.013889 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0241 |  | 0.0061 |  | 0.012 | 0.025 |  | 0.020667
| Median |  | 0.013481 | 0.012 | 0.016 | 0.016 | 0.016 | 0.01 | 0.0245 |  | 0.0061 |  | 0.012 | 0.025 |  | 0.025
| Std Dev. |  | 0.0012659 | 1.9003e-18 | 0 | 0 | 0 | 0 | 0.0012394 |  |  |  | 0 | 3.8006e-18 |  | 0.0075056
| Rel.Std.Dev. |  | 9.114% | 1.584e-14% | 0% | 0% | 0% | 0% | 5.143% |  |  |  | 0% | 1.52e-14% |  | 36.32%
| PDM3 |  | -10.67% | -22.83% | 2.9% | 2.9% | 2.9% | -35.69% | 54.99% |  | -60.77% |  | -22.83% | 60.78% |  | 32.91%
| Z-Score (Absolute) |  | 0.31 | 0 | 0.67 | 0.67 | 0.67 | 0.43 | 2.24 | 1.16 | 1.64 | 1.16 | 0.19 | 2.32 | 1.16 | 0.19 | 1.27
| Analytical results for Mo in OREAS 926b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | < 60 | < 50 | < 50 | [14 empty] | 1
| 1 | 2 | < 60 | < 50 | 60 | [14 empty] | 1
| 1 | 3 | < 60 | < 50 | 70 | [14 empty] | 16
| 1 | 4 | < 60 | < 50 | 70 | [14 empty] | < 50
| 1 | 5 | < 60 | < 50 | 60 | [14 empty] | 7
| 1 | 6 | < 60 | < 50 | 80
| Mean |  |  |  | 68
| Median |  |  |  | 70
| Std Dev. |  |  |  | 8.3666
| Rel.Std.Dev. |  |  |  | 12.3%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 4.16
| Analytical results for Na2O in OREAS 926b (Indicative Value 0.184 wt.%) | [18 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | [15 empty] | 1
| No | No | 12 | 14 | [15 empty] | wt.%
|  |  | BF*XRF | BF*XRF | [15 empty] | 3
3
| 1 | 1 | 0.21 | 0.15702 | [15 empty] | 1
| 1 | 2 | 0.23 | 0.15164 | [15 empty] | 2
| 1 | 3 | 0.21 | 0.1514 | [15 empty] | 16
| 1 | 4 | 0.22 | 0.1512 | [15 empty] | 0.18366
| 1 | 5 | 0.22 | 0.15313 | [15 empty] | 8
| 1 | 6 | 0.2 | 0.14957
| Mean |  | 0.215 | 0.15233
| Median |  | 0.215 | 0.15152
| Std Dev. |  | 0.010488 | 0.0025648
| Rel.Std.Dev. |  | 4.878% | 1.684%
| PDM3 |  | 17.06% | -17.06%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 926b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 05 | 06 | 12 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 100 | < 50 | < 50 | < 35 | [13 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 35 | [13 empty] | 3
| 1 | 3 | < 100 | < 50 | < 50 | < 35 | [13 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 35 | [13 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 35 | [13 empty] | 9
| 1 | 6 | < 100 | < 50 | < 50 | < 35
| Mean |  | 100
| Median |  | 100
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 5.98 | 0 | 0 | 1.35
| Analytical results for Nd in OREAS 926b (Indicative Value < 90 ppm) | [18 empty] | Indicative
Nd | Round | Replicate | Lab | [16 empty] | 1
| No | No | 12 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 0
0
| 1 | 1 | < 90 | [16 empty] | 1
| 1 | 2 | < 90 | [16 empty] | 4
| 1 | 3 | < 90 | [16 empty] | 16
| 1 | 4 | < 90 | [16 empty] | < 90
| 1 | 5 | < 90 | [16 empty] | 10
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 926b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 05 | 06 | 18 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | < 100 | 50 | < 50 | < 80 | [13 empty] | 1
| 1 | 2 | < 100 | 60 | < 50 | < 80 | [13 empty] | 5
| 1 | 3 | < 100 | < 50 | < 50 | < 80 | [13 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 80 | [13 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 80 | [13 empty] | 11
| 1 | 6 | < 100 | 60 | < 50 | < 80
| Mean |  |  | 56.6667
| Median |  |  | 60
| Std Dev. |  |  | 5.7735
| Rel.Std.Dev. |  |  | 10.19%
PDM3
| Z-Score (Absolute) |  | 1.29 | 0.06 | 2.08 | 0.06
| Analytical results for P2O5 in OREAS 926b (Certified Value 0.057 wt.%) | [18 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 04 | 05 | 06 | 12 | 14 | 16 | 18 | [10 empty] | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 3
3
| 1 | 1 | 0.069 | 0.06 | 0.05 | 0.04 | 0.11353 | 0.06 | 0.07 | [10 empty] | 1
| 1 | 2 | 0.069 | 0.06 | 0.05 | 0.036 | 0.11412 | 0.05 | 0.07 | [10 empty] | #N/A
| 1 | 3 | 0.069 | 0.06 | 0.06 | 0.038 | 0.11401 | 0.05 | 0.06 | [10 empty] | 16
| 1 | 4 | 0.069 | 0.06 | 0.05 | 0.04 | 0.11338 | 0.05 | 0.07 | [10 empty] | 0.057097
| 1 | 5 | 0.046 | 0.06 | 0.05 | 0.047 | 0.11309 | 0.05 | 0.07 | [10 empty] | 14
| 1 | 6 | 0.069 | 0.06 | 0.06 | 0.042 | 0.11295 | 0.05 | 0.07
| Mean |  | 0.065167 | 0.06 | 0.053333 | 0.0405 | 0.11351 | 0.051667 | 0.068333
| Median |  | 0.069 | 0.06 | 0.05 | 0.04 | 0.11345 | 0.05 | 0.07
| Std Dev. |  | 0.0093897 | 0 | 0.005164 | 0.0037815 | 0.00047599 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 14.41% | 0% | 9.682% | 9.337% | 0.4193% | 7.902% | 5.974%
| PDM3 |  | 14.13% | 5.085% | -6.592% | -29.07% | 98.81% | -9.511% | 19.68%
| Z-Score (Absolute) |  | 0.4 | 0 | 0.54 | 1.58 | 4.33 | 0.67 | 0.67
| Analytical results for Pb in OREAS 926b (Indicative Value 108 ppm) | [18 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 18 | 19 | 21 |  |  |  |  | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 100 | 200 | < 50 | 80 | 100 | 80 | < 100 | < 90 | < 90 | 100 | 200 | 317.94 | < 100 |  |  |  |  | 1
| 1 | 2 | 100 | 200 | < 50 | 100 | 110 | 50 | 100 | < 90 | < 90 | 100 | < 100 | 325.81 | < 100 |  |  |  |  | 6
| 1 | 3 | 100 | 200 | 50 | 110 | 110 | 70 | < 100 | < 90 | < 90 | 100 | < 100 | 221 | < 100 |  |  |  |  | 16
| 1 | 4 | 200 | 100 | < 50 | 90 | 120 | 50 | < 100 | < 90 | < 90 | 100 | 200 | 335.3 | < 100 |  |  |  |  | 108.1481
| 1 | 5 | 100 | 200 | 60 | 80 | 90 | 50 | < 100 | < 90 | < 90 | < 100 | 100 | 290.66 | < 100 |  |  |  |  | 12
| 1 | 6 | 200 | 100 | 60 | 110 | 120 | 80 | < 100 | < 90 | < 90 | < 100 | 100 | 276.37 | < 100
| Mean |  | 133.3333 | 166.6667 | 56.6667 | 95 | 108.3333 | 63.3333 | 100 |  |  | 100 | 150 | 294.5133
| Median |  | 100 | 200 | 60 | 95 | 110 | 60 | 100 |  |  | 100 | 150 | 304.3
| Std Dev. |  | 51.6398 | 51.6398 | 5.7735 | 13.784 | 11.6905 | 15.0555 |  |  |  | 0 | 57.735 | 42.2849
| Rel.Std.Dev. |  | 38.73% | 30.98% | 10.19% | 14.51% | 10.79% | 23.77% |  |  |  | 0% | 38.49% | 14.36%
| PDM3 |  | 23.29% | 54.11% | -47.6% | -12.16% | 0.1712% | -41.44% | -7.534% |  |  | -7.534% | 38.7% | 172.3%
| Z-Score (Absolute) |  | 1.01 | 1.69 | 0.86 | 0.24 | 0.51 | 0.4 | 0.51 | 0.75 | 0.75 | 0 | 0.67 | 4.27 | 0.67
| Analytical results for Pr in OREAS 926b (Indicative Value 74 ppm) | [18 empty] | Indicative
Pr | Round | Replicate | Lab | [16 empty] | 1
| No | No | 12 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 170 | [16 empty] | 1
| 1 | 2 | 80 | [16 empty] | 7
| 1 | 3 | < 80 | [16 empty] | 16
| 1 | 4 | < 80 | [16 empty] | 74.491
| 1 | 5 | 330 | [16 empty] | 13
| 1 | 6 | < 80
| Mean |  | 193.3333
| Median |  | 170
| Std Dev. |  | 126.6228
| Rel.Std.Dev. |  | 65.49%
| PDM3 |  | 159.5%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 926b (Indicative Value 151 ppm) | [18 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 06 | 12 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 190 | 130 | 90 | [14 empty] | 1
| 1 | 2 | 190 | 120 | 180 | [14 empty] | 8
| 1 | 3 | 190 | 110 | 180 | [14 empty] | 16
| 1 | 4 | 190 | 140 | 90 | [14 empty] | 151.2761
| 1 | 5 | 200 | 120 | 90 | [14 empty] | 14
| 1 | 6 | 200 | 120 | 180
| Mean |  | 193.3333 | 123.3333 | 135
| Median |  | 190 | 120 | 135
| Std Dev. |  | 5.164 | 10.328 | 49.295
| Rel.Std.Dev. |  | 2.671% | 8.374% | 36.51%
| PDM3 |  | 27.8% | -18.47% | -10.76%
| Z-Score (Absolute) |  | 2.74 | 0.67 | 0
| Analytical results for S in OREAS 926b (Certified Value 0.975 wt.%) | [18 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 17 | 18 | 19 | 20 | 21 |  |  |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.96 | 1.02 | 1.01 | 1 | 0.99 | 0.93 | 0.9895 | 0.961 | 1.0834 | 0.99 | 0.99 | 0.836 | 0.052 | 0.91528 |  |  |  | 1
| 1 | 2 | 0.95 | 1.02 | 1.02 | 1 | 0.99 | 0.92 | 0.9803 | 0.969 | 1.0487 | 0.99 | 1.02 | 0.864 | 0.052 | 0.92946 |  |  |  | #N/A
| 1 | 3 | 0.94 | 1.02 | 1.04 | 1 | 0.97 | 0.921 | 0.9887 | 0.985 | 1.0478 | 0.98 | 1.02 | 0.841 | 0.056 | 0.9029 |  |  |  | 16
| 1 | 4 | 0.95 | 1.02 | 1 | 1.01 | 0.98 | 0.926 | 0.9791 | 0.969 | 1.0521 | 0.96 | 1 | 0.884 | 0.052 | 0.91141 |  |  |  | 0.97539
| 1 | 5 | 0.94 | 1.02 | 1.02 | 1.01 | 0.98 | 0.948 | 0.9807 | 0.969 | 1.0557 | 0.97 | 1.03 | 0.889 | 0.052 | 0.90818 |  |  |  | 15
| 1 | 6 | 0.94 | 1.04 | 1.03 | 1.02 | 0.97 | 0.923 | 0.9883 | 1.001 | 1.0409 | 0.96 | 1.03 | 0.839 | 0.052 | 0.90256
| Mean |  | 0.94667 | 1.0233 | 1.02 | 1.0067 | 0.98 | 0.928 | 0.98443 | 0.97567 | 1.0548 | 0.975 | 1.015 | 0.85883 | 0.052667 | 0.91163
| Median |  | 0.945 | 1.02 | 1.02 | 1.005 | 0.98 | 0.9245 | 0.9845 | 0.969 | 1.0504 | 0.975 | 1.02 | 0.8525 | 0.052 | 0.90979
| Std Dev. |  | 0.008165 | 0.008165 | 0.014142 | 0.008165 | 0.0089443 | 0.01045 | 0.004864 | 0.014679 | 0.014869 | 0.013784 | 0.016432 | 0.023676 | 0.001633 | 0.010017
| Rel.Std.Dev. |  | 0.8625% | 0.7979% | 1.386% | 0.8111% | 0.9127% | 1.126% | 0.4941% | 1.504% | 1.41% | 1.414% | 1.619% | 2.757% | 3.101% | 1.099%
| PDM3 |  | -2.945% | 4.915% | 4.573% | 3.206% | 0.4724% | -4.859% | 0.9269% | 0.02815% | 8.137% | -0.04019% | 4.061% | -11.95% | -94.6% | -6.537%
| Z-Score (Absolute) |  | 0.53 | 0.77 | 0.72 | 0.49 | 0.04 | 0.85 | 0.11 | 0.04 | 1.31 | 0.05 | 0.63 | 2.03 | 15.75 | 1.13
| Analytical results for Sb in OREAS 926b (Indicative Value 192 ppm) | [18 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | 100 | 400 | < 50 | [14 empty] | 1
| 1 | 2 | 200 | 390 | < 50 | [14 empty] | 9
| 1 | 3 | < 100 | 450 | 80 | [14 empty] | 16
| 1 | 4 | 100 | 390 | 150 | [14 empty] | 192
| 1 | 5 | 100 | 410 | 790 | [14 empty] | 15
| 1 | 6 | 100 | 400 | < 50
| Mean |  | 120 | 406.6667 | 340
| Median |  | 100 | 400 | 150
| Std Dev. |  | 44.7214 | 22.5093 | 391.28
| Rel.Std.Dev. |  | 37.27% | 5.535% | 115.1%
| PDM3 |  | -37.5% | 111.8% | 77.08%
| Z-Score (Absolute) |  | 0.67 | 2.04 | 0
| Analytical results for SiO2 in OREAS 926b (Certified Value 74.53 wt.%) | [18 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 77.615 | 74.4 | 74.9 | 73.2 | 74.4 | 74.5 | 73.614 | 74.71 | 70.95 | 75.086 | 74.8039 | 75.57 | 74.56 | 75.4 | 48.725 | 73.02 | 72.7609 | 1
| 1 | 2 | 77.519 | 74.9 | 74.8 | 73.8 | 74.6 | 74.6 | 74.197 | 74.83 | 71.66 | 74.658 | 74.9284 | 75.24 | 74.56 | 75.2 | 49.285 | 72.65 | 73.1132 | #N/A
| 1 | 3 | 77.756 | 74.6 | 74.4 | 74.6 | 74.4 | 74 | 74.054 | 74.82 | 71.93 | 74.23 | 74.8902 | 74.61 | 74.66 | 76.2 | 48.745 | 72.3 | 73.0838 | 16
| 1 | 4 | 77.71 | 75.2 | 74.3 | 73.7 | 74.4 | 74.5 | 73.814 | 74.74 | 72.28 | 74.444 | 74.8057 | 75.1 | 74.41 | 76.2 | 49.97 | 72.68 | 73.1266 | 74.5299
| 1 | 5 | 77.722 | 74.8 | 75.2 | 74.8 | 74.2 | 73.9 | 74.039 | 74.72 | 71.11 | 74.872 | 74.8191 | 75.25 | 74.32 | 76.9 | 48.219 | 72.76 | 73.2915 | 16
| 1 | 6 | 77.558 | 75.2 | 75.8 | 74.4 | 74.7 | 74.1 | 73.269 | 74.83 | 72.85 | 74.23 | 74.7768 | 74.16 | 73.59 | 76.4 | 50.076 | 73.92 | 73.1649
| Mean |  | 77.6467 | 74.85 | 74.9 | 74.0833 | 74.45 | 74.2667 | 73.8312 | 74.775 | 71.7967 | 74.5867 | 74.8374 | 74.9883 | 74.35 | 76.05 | 49.17 | 72.8883 | 73.0901
| Median |  | 77.6625 | 74.85 | 74.85 | 74.1 | 74.4 | 74.3 | 73.9265 | 74.78 | 71.795 | 74.551 | 74.8124 | 75.17 | 74.485 | 76.2 | 49.015 | 72.72 | 73.1199
| Std Dev. |  | 0.096751 | 0.32094 | 0.55136 | 0.61455 | 0.17607 | 0.30111 | 0.34356 | 0.057533 | 0.71665 | 0.34946 | 0.058638 | 0.51223 | 0.39151 | 0.63797 | 0.74258 | 0.55578 | 0.17689
| Rel.Std.Dev. |  | 0.1246% | 0.4288% | 0.7361% | 0.8295% | 0.2365% | 0.4054% | 0.4653% | 0.07694% | 0.9982% | 0.4685% | 0.07835% | 0.6831% | 0.5266% | 0.8389% | 1.51% | 0.7625% | 0.242%
| PDM3 |  | 4.182% | 0.4295% | 0.4966% | -0.5992% | -0.1072% | -0.3532% | -0.9375% | 0.3289% | -3.667% | 0.07616% | 0.4125% | 0.6151% | -0.2414% | 2.04% | -34.03% | -2.203% | -1.932%
| Z-Score (Absolute) |  | 4.79 | 0.6 | 0.67 | 0.55 | 0 | 0.27 | 0.93 | 0.49 | 3.98 | 0.2 | 0.58 | 0.81 | 0.15 | 2.4 | 37.88 | 2.34 | 2.04
| Analytical results for Sm in OREAS 926b (Indicative Value < 90 ppm) | [18 empty] | Indicative
Sm | Round | Replicate | Lab | [16 empty] | 1
| No | No | 12 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 0
0
| 1 | 1 | < 90 | [16 empty] | 1
| 1 | 2 | < 90 | [16 empty] | 1
| 1 | 3 | < 90 | [16 empty] | 16
| 1 | 4 | < 90 | [16 empty] | < 90
| 1 | 5 | < 90 | [16 empty] | 7
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 926b (Indicative Value < 50 ppm) | [18 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | [14 empty] | 1
| 1 | 2 | < 100 | 50 | < 50 | [14 empty] | 2
| 1 | 3 | < 100 | 50 | < 50 | [14 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | [14 empty] | < 50
| 1 | 5 | < 100 | 50 | < 50 | [14 empty] | 8
| 1 | 6 | < 100 | 70 | < 50
| Mean |  |  | 55
| Median |  |  | 50
| Std Dev. |  |  | 10
| Rel.Std.Dev. |  |  | 18.18%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 2.7
| Analytical results for Sr in OREAS 926b (Indicative Value 100 ppm) | [18 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 06 | 12 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 100 | < 100 | < 42 | [14 empty] | 1
| 1 | 2 | 100 | < 100 | < 42 | [14 empty] | 3
| 1 | 3 | 100 | < 100 | < 42 | [14 empty] | 16
| 1 | 4 | 100 | < 100 | < 42 | [14 empty] | 100
| 1 | 5 | 100 | < 100 | < 42 | [14 empty] | 9
| 1 | 6 | 100 | < 100 | < 42
| Mean |  | 100
| Median |  | 100
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | 1.17 | 0 | 0.67
| Analytical results for Ta in OREAS 926b (Indicative Value 200 ppm) | [18 empty] | Indicative
Ta | Round | Replicate | Lab | [16 empty] | 1
| No | No | 04 | [16 empty] | ppm
|  |  | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | 300 | [16 empty] | 1
| 1 | 2 | 200 | [16 empty] | 4
| 1 | 3 | 200 | [16 empty] | 16
| 1 | 4 | 200 | [16 empty] | 200
| 1 | 5 | 200 | [16 empty] | 10
| 1 | 6 | 100
| Mean |  | 200
| Median |  | 200
| Std Dev. |  | 63.2456
| Rel.Std.Dev. |  | 31.62%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 926b (Indicative Value < 44 ppm) | [18 empty] | Indicative
Th | Round | Replicate | Lab | [16 empty] | 1
| No | No | 12 | [16 empty] | ppm
|  |  | BF*XRF | [16 empty] | 1
1
| 1 | 1 | < 44 | [16 empty] | 1
| 1 | 2 | < 44 | [16 empty] | 5
| 1 | 3 | < 44 | [16 empty] | 16
| 1 | 4 | < 44 | [16 empty] | < 44
| 1 | 5 | < 44 | [16 empty] | 11
| 1 | 6 | < 44
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 926b (Certified Value 0.477 wt.%) | [18 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 04 | 05 | 06 | 12 | 14 | 16 | 18 | [10 empty] | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 3
3
| 1 | 1 | 0.47 | 0.48 | 0.49 | 0.49 | 0.34833 | 0.48 | 0.47 | [10 empty] | 1
| 1 | 2 | 0.47 | 0.48 | 0.48 | 0.51 | 0.33731 | 0.48 | 0.48 | [10 empty] | #N/A
| 1 | 3 | 0.46 | 0.46 | 0.48 | 0.51 | 0.34326 | 0.47 | 0.48 | [10 empty] | 16
| 1 | 4 | 0.47 | 0.46 | 0.49 | 0.52 | 0.34623 | 0.48 | 0.48 | [10 empty] | 0.47667
| 1 | 5 | 0.47 | 0.47 | 0.48 | 0.51 | 0.34372 | 0.48 | 0.49 | [10 empty] | 17
| 1 | 6 | 0.48 | 0.47 | 0.49 | 0.52 | 0.34968 | 0.47 | 0.49
| Mean |  | 0.47 | 0.47 | 0.485 | 0.51 | 0.34476 | 0.47667 | 0.48167
| Median |  | 0.47 | 0.47 | 0.485 | 0.51 | 0.34498 | 0.48 | 0.48
| Std Dev. |  | 0.0063246 | 0.0089443 | 0.0054772 | 0.010954 | 0.0044255 | 0.005164 | 0.0075277
| Rel.Std.Dev. |  | 1.346% | 1.903% | 1.129% | 2.148% | 1.284% | 1.083% | 1.563%
| PDM3 |  | -1.399% | -1.399% | 1.748% | 6.993% | -27.67% | -2.22e-14% | 1.049%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.84 | 3.37 | 13.34 | 0 | 0.51
| Analytical results for V in OREAS 926b (Indicative Value 197 ppm) | [18 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 05 | 06 | 12 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 200 | 200 | 200 | 202 | [13 empty] | 1
| 1 | 2 | 200 | 200 | 200 | 162 | [13 empty] | 6
| 1 | 3 | 200 | 200 | 200 | 207 | [13 empty] | 16
| 1 | 4 | 200 | 200 | 200 | 196 | [13 empty] | 197.3336
| 1 | 5 | 200 | 200 | 200 | 179 | [13 empty] | 12
| 1 | 6 | 200 | 200 | 200 | 151
| Mean |  | 200 | 200 | 200 | 182.8333
| Median |  | 200 | 200 | 200 | 187.5
| Std Dev. |  | 0 | 0 | 0 | 22.7457
| Rel.Std.Dev. |  | 0% | 0% | 0% | 12.44%
| PDM3 |  | 1.351% | 1.351% | 1.351% | -7.348%
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for W in OREAS 926b (Indicative Value 15.4 ppm) | [18 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 04 | 05 | 06 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [14 empty] | 1
1
| 1 | 1 | < 80 | 10 | 30 | [14 empty] | 1
| 1 | 2 | < 80 | < 10 | 20 | [14 empty] | 7
| 1 | 3 | < 80 | < 10 | 30 | [14 empty] | 16
| 1 | 4 | < 80 | < 10 | 20 | [14 empty] | 15.4167
| 1 | 5 | < 80 | 20 | 10 | [14 empty] | 13
| 1 | 6 | < 80 | 10 | 20
| Mean |  |  | 13.3333 | 21.6667
| Median |  |  | 10 | 20
| Std Dev. |  |  | 5.7735 | 7.5277
| Rel.Std.Dev. |  |  | 43.3% | 34.74%
| PDM3 |  |  | -13.51% | 40.54%
| Z-Score (Absolute) |  | 0.97 | 0.67 | 0
| Analytical results for Y2O3 in OREAS 926b (Indicative Value 67 ppm) | [18 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | [14 empty] | 1
| No | No | 05 | 06 | 12 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 60 | < 50 | < 50 | [14 empty] | 1
| 1 | 2 | 70 | < 50 | < 50 | [14 empty] | 8
| 1 | 3 | 70 | < 50 | < 50 | [14 empty] | 16
| 1 | 4 | 60 | < 50 | < 50 | [14 empty] | 66.6667
| 1 | 5 | 70 | < 50 | < 50 | [14 empty] | 14
| 1 | 6 | 70 | < 50 | < 50
| Mean |  | 66.6667
| Median |  | 70
| Std Dev. |  | 5.164
| Rel.Std.Dev. |  | 7.746%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Zn in OREAS 926b (Certified Value 102 ppm) | [18 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 100 | < 100 | 100 | 80 | 130 | 90 | 120 | 100 | 96 | 105.505 | 200 | 100 | 100 | 188.08 | 160 | 104.8 |  | 1
| 1 | 2 | < 100 | 100 | 100 | 90 | 130 | 100 | 100 | 100 | 80 | 100.983 | 200 | < 100 | 100 | 133.45 | 160 | 97.1 |  | #N/A
| 1 | 3 | < 100 | < 100 | 100 | 100 | 110 | 100 | 120 | 100 | 88 | 101.155 | 200 | < 100 | 100 | 119.1 | 160 | 98.1 |  | 16
| 1 | 4 | < 100 | < 100 | 100 | 90 | 120 | 100 | 110 | 100 | 80 | 104.998 | 200 | < 100 | 100 | 171.95 | 160 | 118.4 |  | 102.0629
| 1 | 5 | < 100 | < 100 | 100 | 110 | 120 | 90 | 120 | 100 | 80 | 100.672 | 200 | < 100 | 100 | 168.81 | 160 | 101.5 |  | 18
| 1 | 6 | < 100 | < 100 | 100 | 100 | 120 | 100 | 100 | 100 | 72 | 102.058 | 200 | < 100 | < 100 | 147.39 | 160 | 103.3
| Mean |  |  | 100 | 100 | 95 | 121.6667 | 96.6667 | 111.6667 | 100 | 82.6667 | 102.5618 | 200 | 100 | 100 | 154.7967 | 160 | 103.8667
| Median |  |  | 100 | 100 | 95 | 120 | 100 | 115 | 100 | 80 | 101.6065 | 200 | 100 | 100 | 158.1 | 160 | 102.4
| Std Dev. |  |  |  | 0 | 10.4881 | 7.5277 | 5.164 | 9.8319 | 0 | 8.2624 | 2.1398 | 0 |  | 0 | 26.0122 | 0 | 7.7078
| Rel.Std.Dev. |  |  |  | 0% | 11.04% | 6.187% | 5.342% | 8.805% | 0% | 9.995% | 2.086% | 0% |  | 0% | 16.8% | 0% | 7.421%
| PDM3 |  |  | -2.021% | -2.021% | -6.92% | 19.21% | -5.287% | 9.41% | -2.021% | -19% | 0.4888% | 95.96% | -2.021% | -2.021% | 51.67% | 56.77% | 1.767%
| Z-Score (Absolute) |  | 2.04 | NA | NA | 0.34 | 0.67 | 0.27 | 0.3 | NA | 0.79 | 0.05 | 3.64 | NA | NA | 1.93 | 2.15 | 0
Results from laboratories 3, 4, 10, 17 and 18 were removed due to their 100 ppm reading resolution.
| Analytical results for Zr in OREAS 926b (Indicative Value 144 ppm) | [18 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 04 | 05 | 06 | 12 | [13 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 100 | 200 | 100 | 1110 | [13 empty] | 1
| 1 | 2 | 100 | 200 | 200 | 1040 | [13 empty] | 9
| 1 | 3 | 100 | 200 | 100 | 1040 | [13 empty] | 16
| 1 | 4 | 100 | 200 | 100 | 1110 | [13 empty] | 144.4444
| 1 | 5 | 100 | 200 | 200 | 1040 | [13 empty] | 15
| 1 | 6 | 100 | 200 | 100 | 1040
| Mean |  | 100 | 200 | 133.3333 | 1063.3333
| Median |  | 100 | 200 | 100 | 1040
| Std Dev. |  | 0 | 0 | 51.6398 | 36.1478
| Rel.Std.Dev. |  | 0% | 0% | 38.73% | 3.399%
| PDM3 |  | -30.77% | 38.46% | -7.692% | 636.2%
| Z-Score (Absolute) |  | 0.9 | 0.45 | 0.45 | 12.06
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C702: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 926b (Certified Value 3.44 wt.%) | [16 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | wt.%
|  |  | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | OxBF*XRF | OxBF*XRF | LOI*Furnace | LOI*Furnace | BF*XRF | LOI*Furnace | 2
3
| 1 | 1 | 3.32 | 3.3 | 3.52 | 3.39 | 3.36 | 3.91 | 3.36 | 3.56 | 2.7033 | 3.3 | 3.42 | 3.45 | 3.79 | 3.58 | 3.53 | 1
| 1 | 2 | 3.29 | 3.34 | 3.39 | 3.32 | 3.37 | 3.97 | 3.36 | 4.01 | 2.689 | 3.3 | 3.45 | 3.4 | 3.78 | 3.6 | 3.5863 | #N/A
| 1 | 3 | 3.24 | 3.29 | 3.48 | 3.31 | 3.35 | 3.94 | 3.28 | 3.65 | 2.7165 | 3.4 | 3.42 | 3.28 | 3.65 | 3.58 | 3.5395 | 16
| 1 | 4 | 3.28 | 3.3 | 3.36 | 3.41 | 3.37 | 3.93 | 3.43 | 3.57 | 2.6933 | 3.3 | 3.43 | 3.41 | 3.66 | 3.58 | 3.5556 | 3.4353
| 1 | 5 | 3.32 | 3.42 | 3.31 | 3.36 | 3.35 | 3.86 | 3.33 | 3.52 | 2.71 | 3.4 | 3.43 | 3.51 | 3.71 | 3.39 | 3.5574 | 20
| 1 | 6 | 3.28 | 3.34 | 3.38 | 3.33 | 3.35 | 4.02 | 3.34 | 3.56 | 2.6765 | 3.2 | 3.42 | 3.38 | 3.71 | 3.58 | 3.5171
| Mean |  | 3.2883 | 3.3317 | 3.4067 | 3.3533 | 3.3583 | 3.9383 | 3.35 | 3.645 | 2.6981 | 3.3167 | 3.4283 | 3.405 | 3.7167 | 3.5517 | 3.5476
| Median |  | 3.285 | 3.32 | 3.385 | 3.345 | 3.355 | 3.935 | 3.35 | 3.565 | 2.6983 | 3.3 | 3.425 | 3.405 | 3.71 | 3.58 | 3.5476
| Std Dev. |  | 0.029944 | 0.048339 | 0.078401 | 0.040332 | 0.0098319 | 0.054191 | 0.04899 | 0.18382 | 0.014694 | 0.075277 | 0.01169 | 0.076616 | 0.058538 | 0.079603 | 0.024345
| Rel.Std.Dev. |  | 0.9106% | 1.451% | 2.301% | 1.203% | 0.2928% | 1.376% | 1.462% | 5.043% | 0.5446% | 2.27% | 0.341% | 2.25% | 1.575% | 2.241% | 0.6862%
| PDM3 |  | -4.278% | -3.016% | -0.8329% | -2.385% | -2.24% | 14.64% | -2.482% | 6.105% | -21.46% | -3.453% | -0.2022% | -0.8815% | 8.191% | 3.388% | 3.271%
| Z-Score (Absolute) |  | 0.89 | 0.56 | 0.01 | 0.39 | 0.36 | 4.07 | 0.42 | 1.83 | 5.4 | 0.67 | 0.18 | 0 | 2.38 | 1.12 | 1.09
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 926b (Certified Value 0.799 wt.%) | [22 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.77 | 0.82712 | 0.79 | 0.84 | 0.75 | 0.83 | 0.79 | 0.81 | 0.86 | 0.78 | 0.77 | 0.84 | 0.77 | 1.324 | 0.76 | 0.75 | 0.83 | 0.863 | 0.75 | 0.813 | 0.8 | 1
| 1 | 2 | 0.78 | 0.8225 | 0.79 | 0.82 | 0.76 | 0.83 | 0.78 | 0.79 | 0.81 | 0.76 | 0.78 | 0.82 | 0.76 | 1.3152 | 0.78 | 0.75 | 0.82 | 0.885 | 0.73 | 0.829 | 0.8 | #N/A
| 1 | 3 | 0.77 | 0.85241 | 0.78 | 0.83 | 0.76 | 0.83 | 0.78 | 0.79 | 0.82 | 0.77 | 0.77 | 0.82 | 0.77 | 1.2878 | 0.78 | 0.75 | 0.82 | 0.873 | 0.75 | 0.821 | 0.8 | 16
| 1 | 4 | 0.78 | 0.85413 | 0.78 | 0.84 | 0.77 | 0.83 | 0.78 | 0.8 | 0.86 | 0.79 | 0.77 | 0.83 | 0.77 | 1.2976 | 0.76 | 0.75 | 0.83 | 0.872 | 0.74 | 0.797 | 0.8 | 0.79916
| 1 | 5 | 0.79 | 0.84202 | 0.79 | 0.84 | 0.76 | 0.83 | 0.79 | 0.8 | 0.83 | 0.87 | 0.77 | 0.83 | 0.77 | 1.2949 | 0.76 | 0.74 | 0.83 | 0.878 | 0.77 | 0.82 | 0.8 | 22
| 1 | 6 | 0.78 | 0.82827 | 0.79 | 0.83 | 0.74 | 0.82 | 0.79 | 0.8 | 0.83 | 0.78 | 0.77 | 0.82 | 0.76 | 1.3097 | 0.76 | 0.76 | 0.83 | 0.888 | 0.75 | 0.839 | 0.8
| Mean |  | 0.77833 | 0.83774 | 0.78667 | 0.83333 | 0.75667 | 0.82833 | 0.785 | 0.79833 | 0.835 | 0.79167 | 0.77167 | 0.82667 | 0.76667 | 1.3049 | 0.76667 | 0.75 | 0.82667 | 0.8765 | 0.74833 | 0.81983 | 0.8
| Median |  | 0.78 | 0.83514 | 0.79 | 0.835 | 0.76 | 0.83 | 0.785 | 0.8 | 0.83 | 0.78 | 0.77 | 0.825 | 0.77 | 1.3036 | 0.76 | 0.75 | 0.83 | 0.8755 | 0.75 | 0.8205 | 0.8
| Std Dev. |  | 0.0075277 | 0.013691 | 0.005164 | 0.008165 | 0.010328 | 0.0040825 | 0.0054772 | 0.0075277 | 0.020736 | 0.039707 | 0.0040825 | 0.008165 | 0.005164 | 0.013689 | 0.010328 | 0.0063246 | 0.005164 | 0.0091815 | 0.013292 | 0.014289 | 1.2162e-16
| Rel.Std.Dev. |  | 0.9672% | 1.634% | 0.6564% | 0.9798% | 1.365% | 0.4929% | 0.6977% | 0.9429% | 2.483% | 5.016% | 0.529% | 0.9877% | 0.6736% | 1.049% | 1.347% | 0.8433% | 0.6247% | 1.048% | 1.776% | 1.743% | 1.52e-14%
| PDM3 |  | -2.606% | 4.828% | -1.564% | 4.276% | -5.317% | 3.65% | -1.772% | -0.1037% | 4.484% | -0.9379% | -3.441% | 3.442% | -4.066% | 63.28% | -4.066% | -6.152% | 3.442% | 9.677% | -6.36% | 2.587% | 0.1049%
| Z-Score (Absolute) |  | 0.35 | 0.91 | 0.18 | 0.82 | 0.82 | 0.71 | 0.21 | 0.07 | 0.85 | 0.07 | 0.5 | 0.67 | 0.6 | 10.86 | 0.6 | 0.96 | 0.67 | 1.74 | 0.99 | 0.53 | NA
Results from laboratory 22 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for S in OREAS 926b (Certified Value 0.986 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.985 | 0.97751 | 0.92 | 0.97 | 0.93 | 0.96 | 0.94 | 0.96 | 1.04 | 0.97 | 1.02 | 1.01 | 0.99 | 0.97899 | 0.99 | 0.97 | 0.98 | 1.02 | 1.05 | 1 | 0.96 | 1
| 1 | 2 | 0.99 | 0.95094 | 0.96 | 0.95 | 0.96 | 0.98 | 0.98 | 0.96 | 1.04 | 0.97 | 1.01 | 1.02 | 0.98 | 0.9789 | 0.99 | 0.98 | 0.97 | 1.02 | 1.02 | 0.99 | 0.95 | #N/A
| 1 | 3 | 0.98 | 0.99948 | 0.94 | 0.96 | 0.97 | 0.95 | 0.95 | 0.97 | 1.04 | 0.98 | 1.02 | 1.05 | 0.98 | 0.97572 | 0.99 | 0.99 | 0.97 | 1.01 | 1.05 | 1 | 0.96 | 16
| 1 | 4 | 1 | 0.99812 | 0.94 | 0.95 | 0.98 | 0.97 | 0.94 | 0.95 | 1.05 | 0.98 | 1.01 | 1.04 | 0.98 | 0.98167 | 1.01 | 0.98 | 0.97 | 1.03 | 1.04 | 1 | 0.96 | 0.98564
| 1 | 5 | 1 | 0.99027 | 0.94 | 0.95 | 0.97 | 0.96 | 0.95 | 0.96 | 1.04 | 1 | 1.01 | 1.05 | 0.97 | 0.97785 | 0.99 | 0.97 | 0.96 | 1.04 | 1.03 | 1 | 0.95 | 23
| 1 | 6 | 0.99 | 1.0008 | 0.95 | 0.94 | 0.96 | 0.96 | 0.95 | 0.95 | 1.06 | 0.99 | 1.02 | 1.06 | 0.98 | 0.97496 | 0.99 | 1 | 0.95 | 1 | 1.05 | 1 | 0.94
| Mean |  | 0.99083 | 0.98619 | 0.94167 | 0.95333 | 0.96167 | 0.96333 | 0.95167 | 0.95833 | 1.045 | 0.98167 | 1.015 | 1.0383 | 0.98 | 0.97802 | 0.99333 | 0.98167 | 0.96667 | 1.02 | 1.04 | 0.99833 | 0.95333
| Median |  | 0.99 | 0.99419 | 0.94 | 0.95 | 0.965 | 0.96 | 0.95 | 0.96 | 1.04 | 0.98 | 1.015 | 1.045 | 0.98 | 0.97838 | 0.99 | 0.98 | 0.97 | 1.02 | 1.045 | 1 | 0.955
| Std Dev. |  | 0.0080104 | 0.019328 | 0.013292 | 0.010328 | 0.017224 | 0.010328 | 0.01472 | 0.0075277 | 0.0083666 | 0.01169 | 0.0054772 | 0.019408 | 0.0063246 | 0.0024384 | 0.008165 | 0.01169 | 0.010328 | 0.014142 | 0.012649 | 0.0040825 | 0.008165
| Rel.Std.Dev. |  | 0.8085% | 1.96% | 1.411% | 1.083% | 1.791% | 1.072% | 1.547% | 0.7855% | 0.8006% | 1.191% | 0.5396% | 1.869% | 0.6454% | 0.2493% | 0.822% | 1.191% | 1.068% | 1.386% | 1.216% | 0.4089% | 0.8565%
| PDM3 |  | 0.5272% | 0.05637% | -4.461% | -3.277% | -2.432% | -2.263% | -3.447% | -2.77% | 6.023% | -0.4028% | 2.979% | 5.346% | -0.5719% | -0.7732% | 0.7809% | -0.4028% | -1.925% | 3.486% | 5.516% | 1.288% | -3.277%
| Z-Score (Absolute) |  | 0.31 | 0.15 | 1.35 | 0.96 | 0.67 | 0.62 | 1.01 | 0.79 | 2.14 | 0 | 1.12 | 1.91 | 0.06 | 0.12 | 0.39 | 0 | 0.51 | 1.29 | 1.97 | 0.56 | 0.96
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 926b (Indicative Value 3.41 ppm) | [19 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 10 | 14 | 15 | 18 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | [14 empty] | 2
2
| 1 | 1 | < 10 | 3.4135 | 5 | < 1 | [14 empty] | 1
| 1 | 2 | < 10 | 3.5632 | < 5 | < 1 | [14 empty] | 11
| 1 | 3 | < 10 | 3.4334 | < 5 | < 1 | [14 empty] | 16
| 1 | 4 | < 10 | 3.3237 | 5 | < 1 | [14 empty] | 3.4075
| 1 | 5 | < 10 | 3.7029 | < 5 | < 1 | [14 empty] | 17
| 1 | 6 | < 10 | 3.4534 | < 5 | < 1
| Mean |  |  | 3.4817 | 5
| Median |  |  | 3.4434 | 5
| Std Dev. |  |  | 0.13293 | 0
| Rel.Std.Dev. |  |  | 3.818% | 0%
| PDM3 |  |  | 2.177% | 46.73%
| Z-Score (Absolute) |  | 1.29 | 0.06 | 0.06 | 2.35
| Analytical results for Al in OREAS 926b (Certified Value 5.78 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 20 | 21 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES |  | 2
3
| 1 | 1 | 5.6818 | 5.795 | 5.901 | 5.822 | 5.769 | 5.822 | 5.71 | 5.89 | 5.69 | 5.5155 | 5.83 | 5.9215 | 5.81 | 5.67 | 5.6617 | 5.92 | 5.5872 |  | 1
| 1 | 2 | 5.992 | 5.795 | 6.166 | 5.927 | 5.769 | 5.848 | 5.85 | 5.89 | 5.76 | 5.7979 | 5.73 | 5.9845 | 5.69 | 5.73 | 5.4648 | 5.83 | 5.7383 |  | #N/A
| 1 | 3 | 5.767 | 5.769 | 6.113 | 6.007 | 5.769 | 5.478 | 5.79 | 5.81 | 5.79 | 6.1355 | 5.77 | 5.9814 | 5.79 | 5.74 | 5.5484 | 5.68 | 5.652 |  | 16
| 1 | 4 | 5.8849 | 5.742 | 6.113 | 6.007 | 5.742 | 5.716 | 5.67 | 5.8 | 5.83 | 5.2702 | 5.75 | 5.9924 | 5.72 | 5.85 | 5.6457 | 5.72 | 5.7437 |  | 5.7814
| 1 | 5 | 5.8181 | 5.742 | 6.113 | 5.875 | 5.822 | 5.663 | 5.8 | 5.83 | 5.82 | 5.2039 | 5.7 | 5.9498 | 5.92 | 5.87 | 5.5959 | 5.74 | 5.49 |  | 25
| 1 | 6 | 5.794 | 5.769 | 6.033 | 5.901 | 5.822 | 5.636 | 5.74 | 5.76 | 5.92 | 5.4746 | 5.81 | 5.9083 | 5.55 | 5.84 | 5.6066 | 5.8 | 5.7275
| Mean |  | 5.823 | 5.7687 | 6.0732 | 5.9232 | 5.7822 | 5.6938 | 5.76 | 5.83 | 5.8017 | 5.5663 | 5.765 | 5.9563 | 5.7467 | 5.7833 | 5.5872 | 5.7817 | 5.6565
| Median |  | 5.806 | 5.769 | 6.113 | 5.914 | 5.769 | 5.6895 | 5.765 | 5.82 | 5.805 | 5.4951 | 5.76 | 5.9656 | 5.755 | 5.79 | 5.6013 | 5.77 | 5.6898
| Std Dev. |  | 0.10615 | 0.023704 | 0.094481 | 0.073644 | 0.032579 | 0.13533 | 0.065727 | 0.051769 | 0.07679 | 0.3486 | 0.048888 | 0.035443 | 0.12533 | 0.080911 | 0.071989 | 0.086814 | 0.10188
| Rel.Std.Dev. |  | 1.823% | 0.4109% | 1.556% | 1.243% | 0.5634% | 2.377% | 1.141% | 0.888% | 1.324% | 6.263% | 0.848% | 0.595% | 2.181% | 1.399% | 1.288% | 1.502% | 1.801%
| PDM3 |  | 0.7184% | -0.2206% | 5.046% | 2.452% | 0.01288% | -1.515% | -0.3705% | 0.8402% | 0.3502% | -3.721% | -0.284% | 3.025% | -0.6012% | 0.03306% | -3.36% | 0.00423% | -2.161%
| Z-Score (Absolute) |  | 0.67 | 0.21 | 4.76 | 2.31 | 0 | 1.44 | 0.35 | 0.79 | 0.33 | 3.52 | 0.27 | 2.85 | 0.57 | 0.03 | 3.18 | 0 | 2.04
| Analytical results for As in OREAS 926b (Certified Value 102 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 200 | 200 | 200 | 200 | 100 | 99 | 101 | 100 | 100 | 110 | 100 | 101 | 99 | 85 | 104.3262 |  |  |  | 1
| 1 | 2 | 200 | 300 | 200 | 100 | 100 | 102 | 98 | 100 | 100 | 100 | 100 | 105 | 103 | 92 | 107.2156 |  |  |  | 1
| 1 | 3 | 100 | 200 | 200 | 200 | 100 | 97 | 103 | 100 | 110 | 100 | < 100 | 102 | 106 | 89 | 107.5425 |  |  |  | 16
| 1 | 4 | 200 | 200 | 200 | 200 | 100 | 99 | 100 | < 100 | 100 | 100 | < 100 | 108 | 102 | 86 | 108.3592 |  |  |  | 102.2085
| 1 | 5 | 200 | 200 | 200 | 200 | 100 | 94 | 102 | < 100 | 100 | 95 | < 100 | 107 | 104 | 88 | 106.1674 |  |  |  | 26
| 1 | 6 | 200 | 200 | 200 | 200 | 100 | 97 | 99 | 100 | 100 | 100 | 100 | 106 | 103 | 86 | 107.1478
| Mean |  | 183.3333 | 216.6667 | 200 | 183.3333 | 100 | 98 | 100.5 | 100 | 101.6667 | 100.8333 | 100 | 104.8333 | 102.8333 | 87.6667 | 106.7931
| Median |  | 200 | 200 | 200 | 200 | 100 | 98 | 100.5 | 100 | 100 | 100 | 100 | 105.5 | 103 | 87 | 107.1817
| Std Dev. |  | 40.8248 | 40.8248 | 0 | 40.8248 | 0 | 2.6833 | 1.8708 | 0 | 4.0825 | 4.916 | 0 | 2.7869 | 2.3166 | 2.582 | 1.3997
| Rel.Std.Dev. |  | 22.27% | 18.84% | 0% | 22.27% | 0% | 2.738% | 1.862% | 0% | 4.016% | 4.875% | 0% | 2.658% | 2.253% | 2.945% | 1.311%
| PDM3 |  | 79.37% | 112% | 95.68% | 79.37% | -2.161% | -4.118% | -1.672% | -2.161% | -0.5302% | -1.345% | -2.161% | 2.568% | 0.6113% | -14.23% | 4.486%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.91 | 0.21 | NA | 0.12 | 0.12 | NA | 1 | 0.44 | 3.79 | 1.55
Results from laboratories 3, 4, 5, 6, 7, 11 and 16 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 926b (Certified Value 131 ppm) | [19 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 17 | 18 | 20 | [11 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [11 empty] | 0
0
| 1 | 1 | 110 | 122 | 135 | 100 | 144 | < 1000 | 245 | [11 empty] | 1
| 1 | 2 | 120 | 122 | 137 | 120 | 143 | < 1000 | 146 | [11 empty] | 16
| 1 | 3 | 120 | 121 | 145 | 100 | 140 | < 1000 | 159 | [11 empty] | 16
| 1 | 4 | 110 | 119 | 135 | 120 | 145 | < 1000 | 154 | [11 empty] | 130.6111
| 1 | 5 | 120 | 123 | 140 | 120 | 146 | < 1000 | 146 | [11 empty] | 27
| 1 | 6 | 110 | 120 | 146 | 120 | 143 | < 1000 | 150
| Mean |  | 115 | 121.1667 | 139.6667 | 113.3333 | 143.5 |  | 166.6667
| Median |  | 115 | 121.5 | 138.5 | 120 | 143.5 |  | 152
| Std Dev. |  | 5.4772 | 1.472 | 4.8854 | 10.328 | 2.0736 |  | 38.6971
| Rel.Std.Dev. |  | 4.763% | 1.215% | 3.498% | 9.113% | 1.445% |  | 23.22%
| PDM3 |  | -11.95% | -7.231% | 6.933% | -13.23% | 9.868% |  | 27.61%
| Z-Score (Absolute) |  | 0.67 | 0.51 | 0 | 0.72 | 0.1 | 9.85 | 0.74
| Analytical results for Ba in OREAS 926b (Certified Value 0.177 wt.%) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | [5 empty] | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | [5 empty] | 3
3
| 1 | 1 | 0.17316 | 0.165 | 0.1736 | 0.18 | 0.177 | 0.188 | 0.20327 | 0.155 | 0.1746 | 0.1684 | 0.1659 | 0.179 | 0.183 | [5 empty] | 1
| 1 | 2 | 0.17963 | 0.172 | 0.1725 | 0.18 | 0.1749 | 0.187 | 0.20095 | 0.161 | 0.1638 | 0.1803 | 0.1616 | 0.177 | 0.189 | [5 empty] | 20
| 1 | 3 | 0.17369 | 0.167 | 0.1749 | 0.18 | 0.1841 | 0.19 | 0.20209 | 0.162 | 0.1665 | 0.1739 | 0.1598 | 0.172 | 0.189 | [5 empty] | 16
| 1 | 4 | 0.17383 | 0.168 | 0.1734 | 0.19 | 0.1704 | 0.19 | 0.20266 | 0.161 | 0.1644 | 0.1787 | 0.163 | 0.174 | 0.19 | [5 empty] | 0.17682
| 1 | 5 | 0.17352 | 0.166 | 0.1739 | 0.18 | 0.1785 | 0.188 | 0.20502 | 0.159 | 0.1582 | 0.1778 | 0.1635 | 0.175 | 0.184 | [5 empty] | 28
| 1 | 6 | 0.17284 | 0.168 | 0.1756 | 0.19 | 0.1869 | 0.19 | 0.20509 | 0.165 | 0.1663 | 0.1791 | 0.1646 | 0.177 | 0.189
| Mean |  | 0.17444 | 0.16767 | 0.17398 | 0.18333 | 0.17863 | 0.18883 | 0.20318 | 0.1605 | 0.16563 | 0.17637 | 0.16307 | 0.17567 | 0.18733
| Median |  | 0.1736 | 0.1675 | 0.17375 | 0.18 | 0.17775 | 0.189 | 0.20297 | 0.161 | 0.16535 | 0.17825 | 0.16325 | 0.176 | 0.189
| Std Dev. |  | 0.0025667 | 0.0024221 | 0.0011089 | 0.005164 | 0.0060437 | 0.0013292 | 0.0016415 | 0.0033317 | 0.0053242 | 0.004472 | 0.002163 | 0.0025033 | 0.0030111
| Rel.Std.Dev. |  | 1.471% | 1.445% | 0.6374% | 2.817% | 3.383% | 0.7039% | 0.8079% | 2.076% | 3.214% | 2.536% | 1.326% | 1.425% | 1.607%
| PDM3 |  | -1.343% | -5.176% | -1.603% | 3.684% | 1.026% | 6.795% | 14.91% | -9.229% | -6.326% | -0.2555% | -7.777% | -0.6514% | 5.947%
| Z-Score (Absolute) |  | 0.1 | 0.67 | 0.14 | 0.65 | 0.25 | 1.11 | 2.32 | 1.28 | 0.85 | 0.06 | 1.06 | 0 | 0.98
| Analytical results for Be in OREAS 926b (Certified Value 2.04 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [7 empty] | 2
3
| 1 | 1 | < 3 | 2 | < 0.2 | 2 | 1.6444 | 2 | 2 | 2.07 | < 5 | < 20 | 2.1281 | [7 empty] | 1
| 1 | 2 | < 3 | 2 | < 0.2 | 2 | 1.8472 | 2 | 2 | 1.93 | < 5 | < 20 | 2.1543 | [7 empty] | #N/A
| 1 | 3 | < 3 | 2 | < 0.2 | 2 | 1.9129 | 2 | 2 | 2 | < 5 | < 20 | 2.1271 | [7 empty] | 16
| 1 | 4 | < 3 | 2 | < 0.2 | 2 | 1.7476 | 2 | 2 | 2.1 | < 5 | < 20 | 2.2391 | [7 empty] | 2.0403
| 1 | 5 | < 3 | 2 | < 0.2 | 2 | 1.7195 | 2 | 2 | 2.06 | < 5 | < 20 | 2.1818 | [7 empty] | 29
| 1 | 6 | < 3 | 2 | 0.4 | 2 | 1.8429 | 2 | 2 | 2.08 | < 5 | < 20 | 2.2489
| Mean |  |  | 2 | 0.4 | 2 | 1.7858 | 2 | 2 | 2.04 |  |  | 2.1799
| Median |  |  | 2 | 0.4 | 2 | 1.7953 | 2 | 2 | 2.065 |  |  | 2.168
| Std Dev. |  |  | 0 |  | 0 | 0.099035 | 0 | 0 | 0.063561 |  |  | 0.053661
| Rel.Std.Dev. |  |  | 0% |  | 0% | 5.546% | 0% | 0% | 3.116% |  |  | 2.462%
| PDM3 |  |  | -1.976% | -80.4% | -1.976% | -12.48% | -1.976% | -1.976% | -0.01522% |  |  | 6.84%
| Z-Score (Absolute) |  | 1.87 | 0 | 6.94 | 0 | 0.8 | 0 | 0 | 0.15 | 1.87 | 29.99 | 0.67
| Analytical results for Bi in OREAS 926b (Certified Value 49.7 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 10 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 45 | 50.5 | 51.5 | 35.6915 | 49 | 50.04 | 50.1 | 49.9 | 52.9186 | [9 empty] | 1
| 1 | 2 | 46 | 50.5 | 51.2 | 36.0132 | 49 | 46.84 | 49.4 | 49.7 | 51.4675 | [9 empty] | 2
| 1 | 3 | 46 | 49.8 | 52.6 | 34.9743 | 47 | 48.14 | 50.1 | 49.1 | 53.676 | [9 empty] | 16
| 1 | 4 | 46 | 49.4 | 52.9 | 35.6 | 47 | 50.39 | 50.8 | 48.5 | 52.2907 | [9 empty] | 49.7053
| 1 | 5 | 46 | 50 | 53.5 | 35.4707 | 47 | 50.42 | 50.4 | 50.2 | 52.664 | [9 empty] | 30
| 1 | 6 | 46 | 49.9 | 53.2 | 35.9216 | 47 | 50.33 | 50.6 | 48.7 | 53.1797
| Mean |  | 45.8333 | 50.0167 | 52.4833 | 35.6118 | 47.6667 | 49.36 | 50.2333 | 49.35 | 52.6994
| Median |  | 46 | 49.95 | 52.75 | 35.6457 | 47 | 50.185 | 50.25 | 49.4 | 52.7913
| Std Dev. |  | 0.40825 | 0.42622 | 0.93256 | 0.37147 | 1.0328 | 1.5117 | 0.49261 | 0.68629 | 0.764
| Rel.Std.Dev. |  | 0.8907% | 0.8522% | 1.777% | 1.043% | 2.167% | 3.063% | 0.9806% | 1.391% | 1.45%
| PDM3 |  | -7.79% | 0.6263% | 5.589% | -28.35% | -4.102% | -0.6948% | 1.062% | -0.7149% | 6.024%
| Z-Score (Absolute) |  | 1.4 | 0.26 | 1.24 | 5.47 | 0.67 | 0 | 0.35 | 0 | 1.33
| Analytical results for Ca in OREAS 926b (Indicative Value 0.067 wt.%) | [19 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.038457 | 0.057 | 0.057 | < 0.05 | < 0.05 | 0.071 | 0.04 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.1 | 0.036 | < 0.1 | < 0.1 | 0.05 | 0.14089 |  | 1
| 1 | 2 | 0.036873 | < 0.05 | 0.071 | 0.093 | < 0.05 | 0.079 | 0.02 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.1 | 0.036 | < 0.1 | < 0.1 | < 0.05 | 0.13574 |  | 12
| 1 | 3 | 0.034192 | 0.071 | 0.071 | 0.093 | < 0.05 | 0.086 | 0.01 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.1 | 0.036 | < 0.1 | < 0.1 | < 0.05 | 0.11472 |  | 16
| 1 | 4 | 0.036662 | 0.071 | 0.079 | 0.129 | < 0.05 | 0.114 | 0.01 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.1 | 0.036 | < 0.1 | < 0.1 | < 0.05 | 0.11773 |  | 0.067418
| 1 | 5 | 0.039237 | 0.071 | 0.071 | 0.172 | < 0.05 | 0.086 | 0.04 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.1 | 0.036 | < 0.1 | < 0.1 | < 0.05 | 0.10808 |  | 18
| 1 | 6 | 0.035282 | 0.057 | 0.071 | < 0.05 | < 0.05 | 0.079 | 0.03 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.1 | 0.036 | < 0.1 | < 0.1 | < 0.05 | 0.11287
| Mean |  | 0.036784 | 0.0654 | 0.07 | 0.12175 |  | 0.085833 | 0.025 | [5 empty] | 0.036 |  |  | 0.05 | 0.12167
| Median |  | 0.036768 | 0.071 | 0.071 | 0.111 |  | 0.0825 | 0.025 | [5 empty] | 0.036 |  |  | 0.05 | 0.11623
| Std Dev. |  | 0.0018875 | 0.0076681 | 0.0071274 | 0.037553 |  | 0.014878 | 0.013784 | [5 empty] | 0 |  |  |  | 0.013366
| Rel.Std.Dev. |  | 5.131% | 11.72% | 10.18% | 30.84% |  | 17.33% | 55.14% | [5 empty] | 0% |  |  |  | 10.99%
| PDM3 |  | -45.44% | -2.993% | 3.83% | 80.59% |  | 27.32% | -62.92% | [5 empty] | -46.6% |  |  | -25.84% | 80.47%
| Z-Score (Absolute) |  | 0.62 | 0.42 | 0.96 | 1.86 | 1.18 | 1.69 | 1.18 | 0 | 1.18 | 0 | 0 | 0 | 0.67 | 0 | 0 | 0.98 | 3.39
| Analytical results for Cd in OREAS 926b (Certified Value < 10 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 08 | 10 | 11 | 13 | 15 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [9 empty] | 1
1
| 1 | 1 | < 100 | < 2 | < 10 | < 50 | < 10 | < 10 | 8.39 | < 10 | < 5 | [9 empty] | 1
| 1 | 2 | < 100 | < 2 | < 10 | < 50 | < 10 | < 10 | 8.67 | < 10 | < 5 | [9 empty] | #N/A
| 1 | 3 | < 100 | < 2 | < 10 | < 50 | < 10 | < 10 | 9.1 | < 10 | < 5 | [9 empty] | 16
| 1 | 4 | < 100 | < 2 | < 10 | < 50 | < 10 | < 10 | 9.34 | < 10 | < 5 | [9 empty] | < 10
| 1 | 5 | < 100 | < 2 | < 10 | < 50 | < 10 | < 10 | 8.87 | < 10 | < 5 | [9 empty] | 31
| 1 | 6 | < 100 | < 2 | < 10 | < 50 | < 10 | < 10 | 8.8 | < 10 | < 5
| Mean | [7 empty] | 8.8617
| Median | [7 empty] | 8.835
| Std Dev. | [7 empty] | 0.33127
| Rel.Std.Dev. | [7 empty] | 3.738%
PDM3
| Z-Score (Absolute) |  | 12.14 | 1.08 | 0 | 5.39 | 0 | 0 | 1.04 | 0 | 0.67
| Analytical results for Ce in OREAS 926b (Certified Value 67 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | [7 empty] | 0
1
| 1 | 1 | 71.3 | 64 | 65.2 | 62.2018 | 76 | 72.1 | 72 | 67.5 | 69.033 | 66.2 | 58.8582 | [7 empty] | 1
| 1 | 2 | 64.6 | 65 | 64.4 | 61.8785 | 76 | 65.8 | 72 | 69 | 71.268 | 67.6 | 60.4239 | [7 empty] | #N/A
| 1 | 3 | 67.2 | 68 | 64.6 | 61.7039 | 73.5 | 66.7 | 72 | 69 | 71.577 | 66 | 61.0129 | [7 empty] | 16
| 1 | 4 | 68.2 | 62 | 65.1 | 61.5232 | 77 | 67.4 | 74 | 72.3 | 70.763 | 67 | 58.9886 | [7 empty] | 67.3343
| 1 | 5 | 70.5 | 63 | 63.2 | 61.665 | 75.5 | 64.1 | 75 | 68.6 | 69.797 | 67.4 | 60.653 | [7 empty] | 32
| 1 | 6 | 70 | 65 | 63.1 | 61.5938 | 74 | 67.9 | 72 | 68.5 | 70.827 | 65.9 | 57.897
| Mean |  | 68.6333 | 64.5 | 64.2667 | 61.761 | 75.3333 | 67.3333 | 72.8333 | 69.15 | 70.5442 | 66.6833 | 59.6389
| Median |  | 69.1 | 64.5 | 64.5 | 61.6845 | 75.75 | 67.05 | 72 | 68.8 | 70.795 | 66.6 | 59.7063
| Std Dev. |  | 2.4889 | 2.0736 | 0.91579 | 0.247 | 1.3292 | 2.6927 | 1.3292 | 1.638 | 0.95503 | 0.74409 | 1.2328
| Rel.Std.Dev. |  | 3.626% | 3.215% | 1.425% | 0.3999% | 1.764% | 3.999% | 1.825% | 2.369% | 1.354% | 1.116% | 2.067%
| PDM3 |  | 1.929% | -4.209% | -4.556% | -8.277% | 11.88% | -0.001461% | 8.167% | 2.697% | 4.767% | -0.9668% | -11.43%
| Z-Score (Absolute) |  | 0.29 | 0.62 | 0.67 | 1.23 | 1.76 | 0 | 1.21 | 0.4 | 0.71 | 0.14 | 1.69
| Analytical results for Co in OREAS 926b (Certified Value 7.24 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS |  |  | 2
3
| 1 | 1 | < 50 | < 50 | < 50 | < 50 | < 50 | 7.7 | 7 | < 10 | 7 | 7.5232 | < 10 | 7.7 | 6.71 | < 10 | 7 | 6.2318 |  |  | 1
| 1 | 2 | < 50 | < 50 | < 50 | < 50 | < 50 | 7.9 | 7 | < 10 | 7 | 7.2398 | < 10 | 7 | 7.48 | < 10 | 7 | 6.2522 |  |  | 4
| 1 | 3 | < 50 | < 50 | < 50 | < 50 | < 50 | 7.7 | 7 | < 10 | 7 | 7.4518 | < 10 | 7.4 | 7.78 | < 10 | 7 | 6.3156 |  |  | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 50 | < 50 | 7.4 | 7 | < 10 | 7 | 7.1937 | < 10 | 7.1 | 7.9 | < 10 | 7 | 6.3604 |  |  | 7.2365
| 1 | 5 | < 50 | < 50 | < 50 | < 50 | < 50 | 7.5 | 7 | < 10 | 7 | 7.5151 | < 10 | 6.8 | 7.82 | < 10 | 7 | 6.274 |  |  | 33
| 1 | 6 | < 50 | < 50 | < 50 | < 50 | < 50 | 7.5 | 7 | < 10 | 7 | 7.5397 | < 10 | 7.2 | 6.88 | < 10 | 7 | 6.4683
| Mean | [6 empty] | 7.6167 | 7 |  | 7 | 7.4106 |  | 7.2 | 7.4283 |  | 7 | 6.3171
| Median | [6 empty] | 7.6 | 7 |  | 7 | 7.4835 |  | 7.15 | 7.63 |  | 7 | 6.2948
| Std Dev. | [6 empty] | 0.18348 | 0 |  | 0 | 0.15375 |  | 0.31623 | 0.51359 |  | 0 | 0.087266
| Rel.Std.Dev. | [6 empty] | 2.409% | 0% |  | 0% | 2.075% |  | 4.392% | 6.914% |  | 0% | 1.381%
| PDM3 | [6 empty] | 5.253% | -3.268% |  | -3.268% | 2.405% |  | -0.5045% | 2.651% |  | -3.268% | -12.71%
| Z-Score (Absolute) |  | 7.25 | 7.25 | 7.25 | 7.25 | 7.25 | 0.13 | 0.13 | 0.94 | 0.13 | 0.04 | 0.94 | 0.04 | 0.05 | 0.94 | 0.13 | 0.4
| Analytical results for Cr in OREAS 926b (Certified Value 84 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 70 | 140 | 140 | 70 | 70 | 110 | 87 | 100 | 84 | 90 | 80 | 82 | 85 | 90 | 100 | 86.2574 |  |  | 1
| 1 | 2 | 70 | 140 | 210 | 70 | 70 | 90 | 94 | 100 | 126 | 80 | 90 | 82 | 93 | 78 | 100 | 99.6252 |  |  | 5
| 1 | 3 | 70 | 140 | 340 | 70 | 70 | 110 | 92 | 100 | 98 | 80 | 80 | 82 | 94 | 90 | 100 | 103.2215 |  |  | 16
| 1 | 4 | 70 | 140 | 480 | 70 | 70 | 80 | 93 | 100 | 111 | 80 | 80 | 82 | 99 | 119 | 100 | 90.0291 |  |  | 83.7043
| 1 | 5 | 70 | 140 | 410 | 70 | 70 | 90 | 92 | < 100 | 88 | 80 | 80 | 89 | 97 | 88 | 100 | 102.6598 |  |  | 34
| 1 | 6 | 70 | 140 | 210 | 70 | 70 | 80 | 92 | 100 | 91 | 70 | 90 | 82 | 94 | 104 | 100 | 88.7427
| Mean |  | 70 | 140 | 298.3333 | 70 | 70 | 93.3333 | 91.6667 | 100 | 99.6667 | 80 | 83.3333 | 83.1667 | 93.6667 | 94.8333 | 100 | 95.0893
| Median |  | 70 | 140 | 275 | 70 | 70 | 90 | 92 | 100 | 94.5 | 80 | 80 | 82 | 94 | 90 | 100 | 94.8271
| Std Dev. |  | 0 | 0 | 132.577 | 0 | 0 | 13.6626 | 2.4221 | 0 | 16.0083 | 6.3246 | 5.164 | 2.8577 | 4.8028 | 14.4557 | 0 | 7.5882
| Rel.Std.Dev. |  | 0% | 0% | 44.44% | 0% | 0% | 14.64% | 2.642% | 0% | 16.06% | 7.906% | 6.197% | 3.436% | 5.128% | 15.24% | 0% | 7.98%
| PDM3 |  | -16.37% | 67.26% | 256.4% | -16.37% | -16.37% | 11.5% | 9.512% | 19.47% | 19.07% | -4.425% | -0.4432% | -0.6423% | 11.9% | 13.3% | 19.47% | 13.6%
| Z-Score (Absolute) |  | 1.63 | 2.79 | 13.13 | 1.63 | 1.63 | NA | 0.13 | NA | 0.39 | NA | NA | 0.67 | 0 | 0.08 | NA | 0.09
Results from laboratories 8, 13, 15 and 20 were removed due to their 10 ppm reading resolution. Results from laboratory 11 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 926b (Certified Value 4.46 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 10 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 4.6 | 4.5 | 4.89 | 4 | 4.4 | 4.6 | 4.3 | 4.6 | 4.2092 | [9 empty] | 1
| 1 | 2 | 4.4 | 4.2 | 4.76 | 4 | 4.2 | 4.8 | 4.2 | 5 | 4.2738 | [9 empty] | #N/A
| 1 | 3 | 4.3 | 4.5 | 4.76 | 4 | 4.1 | 4.49 | 4.4 | 5.8 | 4.3776 | [9 empty] | 16
| 1 | 4 | 4.7 | 4.5 | 4.73 | 4 | 4.1 | 4.89 | 4.5 | 7.3 | 4.3422 | [9 empty] | 4.456
| 1 | 5 | 4.4 | 4.7 | 4.64 | 4 | 4 | 4.52 | 4.5 | 5.9 | 4.3574 | [9 empty] | 35
| 1 | 6 | 4.3 | 4.6 | 4.62 | 4 | 4.2 | 4.67 | 4.3 | 5.7 | 4.3233
| Mean |  | 4.45 | 4.5 | 4.7333 | 4 | 4.1667 | 4.6617 | 4.3667 | 5.7167 | 4.3139
| Median |  | 4.4 | 4.5 | 4.745 | 4 | 4.15 | 4.635 | 4.35 | 5.75 | 4.3328
| Std Dev. |  | 0.16432 | 0.16733 | 0.097502 | 0 | 0.13663 | 0.15791 | 0.12111 | 0.92826 | 0.062297
| Rel.Std.Dev. |  | 3.693% | 3.718% | 2.06% | 0% | 3.279% | 3.387% | 2.773% | 16.24% | 1.444%
| PDM3 |  | -0.1354% | 0.9866% | 6.223% | -10.23% | -6.494% | 4.615% | -2.006% | 28.29% | -3.189%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 1 | NA | 1.2 | 0.72 | 0.42 | 4.82 | 0.62
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 926b (Certified Value 0.865 wt.%) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | 3
3
| 1 | 1 | 0.88104 | 0.86 | 0.85 | 0.88 | 0.86 | 0.81 | 0.93 | 0.8729 | 0.852 | 0.8675 | 0.871 | 0.83252 | 0.88 | 0.852 | 0.875 | 0.8216 | 0.858 | 0.87567 | 1
| 1 | 2 | 0.87288 | 0.82 | 0.87 | 0.89 | 0.84 | 0.77 | 0.915 | 0.8652 | 0.85 | 0.8705 | 0.867 | 0.83132 | 0.87 | 0.835 | 0.8808 | 0.789 | 0.858 | 0.86557 | #N/A
| 1 | 3 | 0.88154 | 0.89 | 0.87 | 0.89 | 0.85 | 0.78 | 0.923 | 0.8619 | 0.862 | 0.917 | 0.867 | 0.83588 | 0.875 | 0.843 | 0.8853 | 0.7967 | 0.839 | 0.93829 | 16
| 1 | 4 | 0.88434 | 0.85 | 0.86 | 0.89 | 0.85 | 0.82 | 0.949 | 0.8591 | 0.872 | 0.8411 | 0.865 | 0.83564 | 0.875 | 0.836 | 0.8858 | 0.8161 | 0.846 | 0.89385 | 0.86478
| 1 | 5 | 0.87972 | 0.84 | 0.88 | 0.89 | 0.86 | 0.8 | 0.922 | 0.8608 | 0.869 | 0.8724 | 0.872 | 0.83516 | 0.86 | 0.874 | 0.8632 | 0.8125 | 0.855 | 0.86911 | 36
| 1 | 6 | 0.87755 | 0.86 | 0.87 | 0.87 | 0.87 | 0.79 | 0.929 | 0.8495 | 0.9 | 0.8946 | 0.872 | 0.83364 | 0.88 | 0.855 | 0.8614 | 0.8229 | 0.856 | 0.86153
| Mean |  | 0.87951 | 0.85333 | 0.86667 | 0.885 | 0.855 | 0.795 | 0.928 | 0.86157 | 0.8675 | 0.87718 | 0.869 | 0.83403 | 0.87333 | 0.84917 | 0.87525 | 0.8098 | 0.852 | 0.884
| Median |  | 0.88038 | 0.855 | 0.87 | 0.89 | 0.855 | 0.795 | 0.926 | 0.86135 | 0.8655 | 0.87145 | 0.869 | 0.8344 | 0.875 | 0.8475 | 0.8779 | 0.8143 | 0.8555 | 0.87239
| Std Dev. |  | 0.0039404 | 0.023381 | 0.010328 | 0.0083666 | 0.010488 | 0.018708 | 0.011628 | 0.0076686 | 0.018196 | 0.025892 | 0.0030332 | 0.0018476 | 0.0075277 | 0.014634 | 0.010772 | 0.013871 | 0.0077717 | 0.028914
| Rel.Std.Dev. |  | 0.448% | 2.74% | 1.192% | 0.9454% | 1.227% | 2.353% | 1.253% | 0.8901% | 2.098% | 2.952% | 0.349% | 0.2215% | 0.862% | 1.723% | 1.231% | 1.713% | 0.9122% | 3.271%
| PDM3 |  | 1.704% | -1.324% | 0.2183% | 2.338% | -1.131% | -8.069% | 7.311% | -0.3715% | 0.3147% | 1.434% | 0.4881% | -3.556% | 0.9892% | -1.805% | 1.211% | -6.358% | -1.478% | 2.223%
| Z-Score (Absolute) |  | 0.64 | 0.71 | 0.02 | 0.92 | 0.62 | 3.71 | 3.14 | 0.28 | 0.02 | 0.52 | 0.1 | 1.7 | 0.32 | 0.92 | 0.42 | 2.95 | 0.78 | 0.87
| Analytical results for Dy in OREAS 926b (Certified Value 3.9 ppm) | [19 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 4.2 | 4 | 4.3 | 3.127 | 3.5 | 4.15 | 3.7 | 4.14 | 3.557 | 4.13 | [8 empty] | 1
| 1 | 2 | 4.3 | 4 | 4.1 | 3.1672 | 4 | 3.94 | 4.07 | 4.1 | 3.658 | 4.24 | [8 empty] | #N/A
| 1 | 3 | 4.3 | 4 | 4.3 | 3.1688 | 3.5 | 3.95 | 4.01 | 4.33 | 3.704 | 4 | [8 empty] | 16
| 1 | 4 | 4.3 | 4 | 4.3 | 3.1358 | 4 | 3.92 | 3.63 | 4.07 | 3.625 | 4.17 | [8 empty] | 3.8955
| 1 | 5 | 4.2 | 4 | 4.3 | 3.1829 | 3.5 | 3.74 | 3.72 | 4.15 | 3.647 | 4.14 | [8 empty] | 37
| 1 | 6 | 4.2 | 4 | 4.3 | 3.1292 | 3.5 | 3.93 | 3.97 | 4.22 | 3.536 | 4.2
| Mean |  | 4.25 | 4 | 4.2667 | 3.1518 | 3.6667 | 3.9383 | 3.85 | 4.1683 | 3.6212 | 4.1467
| Median |  | 4.25 | 4 | 4.3 | 3.1515 | 3.5 | 3.935 | 3.845 | 4.145 | 3.636 | 4.155
| Std Dev. |  | 0.054772 | 0 | 0.08165 | 0.023972 | 0.2582 | 0.13014 | 0.18772 | 0.09411 | 0.063672 | 0.082381
| Rel.Std.Dev. |  | 1.289% | 0% | 1.914% | 0.7606% | 7.042% | 3.304% | 4.876% | 2.258% | 1.758% | 1.987%
| PDM3 |  | 9.1% | 2.682% | 9.528% | -19.09% | -5.875% | 1.099% | -1.168% | 7.003% | -7.043% | 6.447%
| Z-Score (Absolute) |  | 0.77 | NA | 0.81 | 1.95 | 0.67 | 0 | 0.22 | 0.57 | 0.79 | 0.52
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 926b (Certified Value 2.4 ppm) | [19 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 2.5 | 2 | 2.1 | 1.6305 | 2 | 2.35 | 1.96 | 2.44 | 2.488 | 2.42 | [8 empty] | 1
| 1 | 2 | 2.3 | 2 | 2.3 | 1.7592 | 2.5 | 2.22 | 2.33 | 2.29 | 2.481 | 2.5 | [8 empty] | 21
| 1 | 3 | 2.4 | 2 | 2.2 | 1.6348 | 2.5 | 2.25 | 2.17 | 2.57 | 2.536 | 2.45 | [8 empty] | 16
| 1 | 4 | 2.6 | 2 | 2.4 | 1.7071 | 2.5 | 2.2 | 2.34 | 2.47 | 2.513 | 2.31 | [8 empty] | 2.3956
| 1 | 5 | 2.5 | 2 | 2.3 | 1.7891 | 2.5 | 2.19 | 2.43 | 2.41 | 2.54 | 2.48 | [8 empty] | 38
| 1 | 6 | 2.4 | 2 | 2.4 | 1.5829 | 2.5 | 2.25 | 2.32 | 2.41 | 2.494 | 2.42
| Mean |  | 2.45 | 2 | 2.2833 | 1.684 | 2.4167 | 2.2433 | 2.2583 | 2.4317 | 2.5087 | 2.43
| Median |  | 2.45 | 2 | 2.3 | 1.671 | 2.5 | 2.235 | 2.325 | 2.425 | 2.5035 | 2.435
| Std Dev. |  | 0.10488 | 0 | 0.1169 | 0.080922 | 0.20412 | 0.05785 | 0.16845 | 0.091305 | 0.025121 | 0.066933
| Rel.Std.Dev. |  | 4.281% | 0% | 5.12% | 4.805% | 8.447% | 2.579% | 7.459% | 3.755% | 1.001% | 2.754%
| PDM3 |  | 2.27% | -16.51% | -4.687% | -29.71% | 0.8783% | -6.357% | -5.731% | 1.504% | 4.719% | 1.435%
| Z-Score (Absolute) |  | 0.24 | NA | 0.98 | 5.37 | 0 | 1.27 | 1.16 | 0.11 | 0.67 | 0.1
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 926b (Certified Value 1.1 ppm) | [19 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 1.3 | < 2 | 1.1 | 1.4 | 1.15 | 1.13 | 1.01 | 1.08 | [10 empty] | 1
| 1 | 2 | 1.1 | < 2 | 1.1 | 1.2 | 1.12 | 1.09 | 1.01 | 1.04 | [10 empty] | 22
| 1 | 3 | 1.2 | < 2 | 1.1 | 1.2 | 1.08 | 1.1 | 1.05 | 1.11 | [10 empty] | 16
| 1 | 4 | 1.2 | < 2 | 1.1 | 1.8 | 1.09 | 1.19 | 1.12 | 1.05 | [10 empty] | 1.1025
| 1 | 5 | 1.1 | < 2 | 1.1 | 1.4 | 1.05 | 1.17 | 1 | 1.14 | [10 empty] | 39
| 1 | 6 | 1.1 | < 2 | 1.1 | 1.4 | 1.1 | 1.21 | 1.1 | 1.06
| Mean |  | 1.1667 |  | 1.1 | 1.4 | 1.0983 | 1.1483 | 1.0483 | 1.08
| Median |  | 1.15 |  | 1.1 | 1.4 | 1.095 | 1.15 | 1.03 | 1.07
| Std Dev. |  | 0.08165 |  | 2.4324e-16 | 0.21909 | 0.034303 | 0.04916 | 0.051153 | 0.038471
| Rel.Std.Dev. |  | 6.999% |  | 2.211e-14% | 15.65% | 3.123% | 4.281% | 4.879% | 3.562%
| PDM3 |  | 5.82% |  | -0.2268% | 26.98% | -0.3779% | 4.157% | -4.913% | -2.041%
| Z-Score (Absolute) |  | 0.91 | 1.34 | 0.01 | 4.06 | 0.01 | 0.66 | 0.69 | 0.26
| Analytical results for Fe in OREAS 926b (Certified Value 3.69 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | 3.6608 | 3.651 | 3.609 | 3.721 | 3.693 | 3.679 | 3.77 | 3.72 | 3.73 | 3.4177 | 3.93 | 3.7145 | 3.5 | 3.707 | 3.79 | 3.6548 | 3.65 | 3.7346 | 1
| 1 | 2 | 3.7335 | 3.588 | 3.63 | 3.791 | 3.651 | 3.658 | 3.75 | 3.71 | 3.71 | 3.4798 | 3.94 | 3.7192 | 3.57 | 3.693 | 3.84 | 3.6584 | 3.62 | 3.8886 | #N/A
| 1 | 3 | 3.7149 | 3.763 | 3.693 | 3.742 | 3.651 | 3.483 | 3.77 | 3.66 | 3.81 | 3.678 | 4 | 3.7472 | 3.45 | 3.756 | 3.81 | 3.665 | 3.56 | 3.8584 | 16
| 1 | 4 | 3.7281 | 3.637 | 3.672 | 3.749 | 3.637 | 3.7 | 3.77 | 3.68 | 3.86 | 3.308 | 3.95 | 3.7298 | 3.46 | 3.714 | 3.91 | 3.7489 | 3.58 | 3.8761 | 3.6903
| 1 | 5 | 3.6635 | 3.567 | 3.693 | 3.728 | 3.679 | 3.602 | 3.76 | 3.68 | 3.87 | 3.4467 | 3.92 | 3.7167 | 3.6 | 3.7 | 3.91 | 3.6797 | 3.62 | 3.5922 | 40
| 1 | 6 | 3.6963 | 3.658 | 3.693 | 3.791 | 3.7 | 3.588 | 3.75 | 3.64 | 3.9 | 3.5754 | 3.99 | 3.7557 | 3.6 | 3.672 | 3.91 | 3.7031 | 3.62 | 3.7669
| Mean |  | 3.6995 | 3.644 | 3.665 | 3.7537 | 3.6685 | 3.6183 | 3.7617 | 3.6817 | 3.8133 | 3.4843 | 3.955 | 3.7305 | 3.53 | 3.707 | 3.8617 | 3.685 | 3.6083 | 3.7861
| Median |  | 3.7056 | 3.644 | 3.6825 | 3.7455 | 3.665 | 3.63 | 3.765 | 3.68 | 3.835 | 3.4633 | 3.945 | 3.7245 | 3.535 | 3.7035 | 3.875 | 3.6724 | 3.62 | 3.8126
| Std Dev. |  | 0.031692 | 0.068586 | 0.036775 | 0.030566 | 0.02572 | 0.079278 | 0.0098319 | 0.029944 | 0.078145 | 0.12866 | 0.032711 | 0.017259 | 0.068702 | 0.028 | 0.055287 | 0.03594 | 0.032506 | 0.11354
| Rel.Std.Dev. |  | 0.8567% | 1.882% | 1.003% | 0.8143% | 0.7011% | 2.191% | 0.2614% | 0.8133% | 2.049% | 3.693% | 0.8271% | 0.4627% | 1.946% | 0.7553% | 1.432% | 0.9753% | 0.9009% | 2.999%
| PDM3 |  | 0.2507% | -1.254% | -0.6845% | 1.718% | -0.5896% | -1.949% | 1.935% | -0.2328% | 3.335% | -5.582% | 7.174% | 1.091% | -4.343% | 0.4537% | 4.645% | -0.1429% | -2.22% | 2.598%
| Z-Score (Absolute) |  | 0.07 | 0.5 | 0.28 | 0.63 | 0.24 | 0.76 | 0.72 | 0.11 | 1.25 | 2.14 | 2.71 | 0.39 | 1.67 | 0.15 | 1.75 | 0.07 | 0.86 | 0.97
| Analytical results for Ga in OREAS 926b (Certified Value 16.9 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 10 | 13 | 14 | 16 | 17 | 18 | 20 | 21 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 16.1 | 17 | 16 | 14.461 | 18.8 | 17.87 | 15 | 18 | 17.0849 | [9 empty] | 1
| 1 | 2 | 17.3 | 17 | 16 | 14.451 | 17.3 | 18.05 | 16 | 17 | 16.818 | [9 empty] | #N/A
| 1 | 3 | 17 | 17 | 16 | 14.4437 | 17.8 | 18.1 | 17 | 17.3 | 16.8867 | [9 empty] | 16
| 1 | 4 | 17.1 | 16 | 17 | 14.4137 | 17.2 | 18.88 | 17 | 16.9 | 17.0166 | [9 empty] | 16.8884
| 1 | 5 | 16.7 | 17 | 17 | 14.5988 | 16.7 | 18.26 | 17 | 17.2 | 16.8219 | [9 empty] | 41
| 1 | 6 | 16.8 | 17 | 16 | 14.7127 | 17.4 | 18.45 | 17 | 16.8 | 17.0057
| Mean |  | 16.8333 | 16.8333 | 16.3333 | 14.5135 | 17.5333 | 18.2683 | 16.5 | 17.2 | 16.939
| Median |  | 16.9 | 17 | 16 | 14.456 | 17.35 | 18.18 | 17 | 17.1 | 16.9462
| Std Dev. |  | 0.41793 | 0.40825 | 0.5164 | 0.11701 | 0.71461 | 0.3583 | 0.83666 | 0.43359 | 0.1121
| Rel.Std.Dev. |  | 2.483% | 2.425% | 3.162% | 0.8062% | 4.076% | 1.961% | 5.071% | 2.521% | 0.6618%
| PDM3 |  | -0.3262% | -0.3262% | -3.287% | -14.06% | 3.819% | 8.171% | -2.3% | 1.845% | 0.2992%
| Z-Score (Absolute) |  | 0 | 0 | 0.92 | 4.27 | 1.29 | 2.64 | 0.61 | 0.67 | 0.19
| Analytical results for Gd in OREAS 926b (Certified Value 4.69 ppm) | [19 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 4.9 | < 5 | 4.4 | 4 | 4.91 | 4.12 | 4.37 | 4.711 | 4.81 | [9 empty] | 1
| 1 | 2 | 4.5 | < 5 | 4.5 | 4 | 4.45 | 4.16 | 4.58 | 4.935 | 4.84 | [9 empty] | 6
| 1 | 3 | 4.8 | 5 | 4.6 | 4 | 4.55 | 4.3 | 4.6 | 4.96 | 5.15 | [9 empty] | 16
| 1 | 4 | 4.7 | < 5 | 4.4 | 4 | 4.62 | 4.17 | 4.73 | 4.992 | 4.83 | [9 empty] | 4.6863
| 1 | 5 | 4.7 | < 5 | 4.4 | 4 | 4.36 | 4.82 | 4.77 | 4.99 | 4.85 | [9 empty] | 42
| 1 | 6 | 4.9 | 5 | 4.4 | 4 | 4.62 | 4.21 | 4.57 | 4.906 | 4.86
| Mean |  | 4.75 | 5 | 4.45 | 4 | 4.585 | 4.2967 | 4.6033 | 4.9157 | 4.89
| Median |  | 4.75 | 5 | 4.4 | 4 | 4.585 | 4.19 | 4.59 | 4.9475 | 4.845
| Std Dev. |  | 0.15166 | 0 | 0.083666 | 0 | 0.18876 | 0.26357 | 0.14137 | 0.10552 | 0.12853
| Rel.Std.Dev. |  | 3.193% | 0% | 1.88% | 0% | 4.117% | 6.134% | 3.071% | 2.147% | 2.628%
| PDM3 |  | 1.359% | 6.693% | -5.043% | -14.65% | -2.162% | -8.315% | -1.771% | 4.894% | 4.346%
| Z-Score (Absolute) |  | 0.47 | 3.76 | 0.43 | NA | 0.03 | 0.89 | 0.03 | 0.96 | 0.88
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 926b (Certified Value 2.08 ppm) | [19 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 08 | 10 | 13 | 17 | 18 | 20 | 21 | [11 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [11 empty] | 2
3
| 1 | 1 | 2.2 | 2 | 2 | 2.49 | 2 | 2 | 2.0686 | [11 empty] | 1
| 1 | 2 | 2.3 | 2 | 2 | 2.29 | 2 | 2 | 2.1692 | [11 empty] | 17
| 1 | 3 | 2.2 | 2 | 2 | 2.19 | 2 | 2 | 2.1233 | [11 empty] | 16
| 1 | 4 | 2 | 2 | 2 | 2.03 | 2 | 2 | 2.1319 | [11 empty] | 2.0751
| 1 | 5 | 2.1 | 2 | 2 | 2.18 | 2 | 2 | 2.1523 | [11 empty] | 43
| 1 | 6 | 2.6 | 2 | 2 | 2.11 | 2 | 2 | 2.2591
| Mean |  | 2.2333 | 2 | 2 | 2.215 | 2 | 2 | 2.1507
| Median |  | 2.2 | 2 | 2 | 2.185 | 2 | 2 | 2.1421
| Std Dev. |  | 0.20656 | 0 | 0 | 0.16022 | 0 | 0 | 0.063151
| Rel.Std.Dev. |  | 9.249% | 0% | 0% | 7.233% | 0% | 0% | 2.936%
| PDM3 |  | 7.625% | -3.619% | -3.619% | 6.742% | -3.619% | -3.619% | 3.645%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Hf in OREAS 926b (Indicative Value 4.16 ppm) | [19 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 08 | 15 | 16 | 17 | 18 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 10 | 4 | 4.5 | 4.87 | 4 | [13 empty] | 1
| 1 | 2 | < 10 | 4 | 4.1 | 4.46 | 4 | [13 empty] | 13
| 1 | 3 | < 10 | 4 | 4.1 | 4.24 | 4 | [13 empty] | 16
| 1 | 4 | < 10 | 6 | 4.1 | 4.66 | 4 | [13 empty] | 4.1571
| 1 | 5 | < 10 | 4 | 4 | 4.42 | 4 | [13 empty] | 19
| 1 | 6 | < 10 | 4 | 4.2 | 4.52 | 4
| Mean |  |  | 4.3333 | 4.1667 | 4.5283 | 4
| Median |  |  | 4 | 4.1 | 4.49 | 4
| Std Dev. |  |  | 0.8165 | 0.17512 | 0.21619 | 0
| Rel.Std.Dev. |  |  | 18.84% | 4.203% | 4.774% | 0%
| PDM3 |  |  | 4.24% | 0.2305% | 8.931% | -3.779%
| Z-Score (Absolute) |  | 2.31 | 0 | 0.58 | 0.67 | 1.15
| Analytical results for Ho in OREAS 926b (Certified Value 0.81 ppm) | [19 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.9 | < 2 | 0.8 | 0.8 | 0.81 | 0.75 | 0.86 | 0.813 | 0.75 | [9 empty] | 1
| 1 | 2 | 0.9 | < 2 | 0.8 | 0.8 | 0.74 | 0.81 | 0.82 | 0.817 | 0.83 | [9 empty] | 23
| 1 | 3 | 0.9 | < 2 | 0.8 | 0.8 | 0.77 | 0.82 | 0.85 | 0.828 | 0.82 | [9 empty] | 16
| 1 | 4 | 0.9 | < 2 | 0.8 | 1.2 | 0.77 | 0.84 | 0.82 | 0.819 | 0.79 | [9 empty] | 0.80529
| 1 | 5 | 0.9 | < 2 | 0.8 | 0.8 | 0.72 | 0.84 | 0.85 | 0.83 | 0.8 | [9 empty] | 44
| 1 | 6 | 0.9 | < 2 | 0.8 | 0.8 | 0.78 | 0.84 | 0.83 | 0.815 | 0.79
| Mean |  | 0.9 |  | 0.8 | 0.86667 | 0.765 | 0.81667 | 0.83833 | 0.82033 | 0.79667
| Median |  | 0.9 |  | 0.8 | 0.8 | 0.77 | 0.83 | 0.84 | 0.818 | 0.795
| Std Dev. |  | 0 |  | 1.2162e-16 | 0.1633 | 0.031464 | 0.035024 | 0.017224 | 0.0070333 | 0.028048
| Rel.Std.Dev. |  | 0% |  | 1.52e-14% | 18.84% | 4.113% | 4.289% | 2.055% | 0.8574% | 3.521%
| PDM3 |  | 11.76% |  | -0.6564% | 7.622% | -5.003% | 1.413% | 4.104% | 1.869% | -1.07%
| Z-Score (Absolute) |  | 2.27 | 5.12 | 0.58 | 1.32 | 1.58 | 0.1 | 0.51 | 0 | 0.67
| Analytical results for In in OREAS 926b (Certified Value 0.8 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 10 | 13 | 15 | 17 | 18 | 20 | 21 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 0.7 | 0.8 | 0.9 | 0.8 | 0.85 | 0.8 | 0.8 | 0.8253 | [10 empty] | 1
| 1 | 2 | 0.8 | 0.9 | 0.9 | 0.8 | 0.8 | 0.8 | 0.8 | 0.80743 | [10 empty] | #N/A
| 1 | 3 | 0.9 | 0.8 | 0.9 | 0.6 | 0.95 | 0.8 | 0.7 | 0.80026 | [10 empty] | 16
| 1 | 4 | 0.8 | 0.8 | 0.9 | 0.8 | 0.8 | 0.8 | 0.8 | 0.83961 | [10 empty] | 0.802
| 1 | 5 | 0.8 | 0.8 | 0.9 | 0.8 | 0.81 | 0.8 | 0.8 | 0.79752 | [10 empty] | 45
| 1 | 6 | 0.9 | 0.7 | 0.9 | 0.8 | 0.81 | 0.8 | 0.7 | 0.82973
| Mean |  | 0.81667 | 0.8 | 0.9 | 0.76667 | 0.83667 | 0.8 | 0.76667 | 0.81664
| Median |  | 0.8 | 0.8 | 0.9 | 0.8 | 0.81 | 0.8 | 0.8 | 0.81637
| Std Dev. |  | 0.075277 | 0.063246 | 0 | 0.08165 | 0.058538 | 1.2162e-16 | 0.05164 | 0.017278
| Rel.Std.Dev. |  | 9.218% | 7.906% | 0% | 10.65% | 6.997% | 1.52e-14% | 6.736% | 2.116%
| PDM3 |  | 1.829% | -0.249% | 12.22% | -4.405% | 4.323% | -0.249% | -4.405% | 1.826%
| Z-Score (Absolute) |  | 0.31 | 0.31 | 3.37 | 1.53 | 1.04 | 0.31 | 1.53 | 0.31
| Analytical results for K in OREAS 926b (Certified Value 2.46 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES |  | 2
3
| 1 | 1 | 2.4409 | 2.507 | 2.449 | 2.549 | 2.441 | 2.49 | 2.5 | 2.5 | 2.4 | 2.2584 | 2.52 | 2.2 | 2.441 | 2.53 | 2.4112 | 2.41 | 1.5223 |  | 1
| 1 | 2 | 2.3889 | 2.399 | 2.557 | 2.623 | 2.407 | 2.391 | 2.4 | 2.5 | 2.4 | 2.2688 | 2.48 | 2.5 | 2.449 | 2.5 | 2.362 | 2.34 | 1.5259 |  | #N/A
| 1 | 3 | 2.4309 | 2.524 | 2.573 | 2.607 | 2.399 | 2.316 | 2.5 | 2.46 | 2.5 | 2.4359 | 2.5 | 2.5 | 2.474 | 2.53 | 2.3508 | 2.3 | 1.7066 |  | 16
| 1 | 4 | 2.444 | 2.549 | 2.573 | 2.656 | 2.366 | 2.441 | 2.4 | 2.46 | 2.5 | 2.1401 | 2.5 | 2.5 | 2.441 | 2.57 | 2.4302 | 2.3 | 1.5777 |  | 2.4609
| 1 | 5 | 2.3927 | 2.507 | 2.582 | 2.681 | 2.432 | 2.499 | 2.5 | 2.47 | 2.4 | 2.344 | 2.47 | 2.4 | 2.424 | 2.58 | 2.4079 | 2.31 | 1.488 |  | 46
| 1 | 6 | 2.4373 | 2.524 | 2.524 | 2.507 | 2.482 | 2.441 | 2.5 | 2.44 | 2.5 | 2.4855 | 2.48 | 2.6 | 2.449 | 2.53 | 2.4123 | 2.33 | 1.744
| Mean |  | 2.4224 | 2.5017 | 2.543 | 2.6038 | 2.4212 | 2.4297 | 2.4667 | 2.4717 | 2.45 | 2.3221 | 2.4917 | 2.45 | 2.4463 | 2.54 | 2.3957 | 2.3317 | 1.5941
| Median |  | 2.4341 | 2.5155 | 2.565 | 2.615 | 2.4195 | 2.441 | 2.5 | 2.465 | 2.45 | 2.3064 | 2.49 | 2.5 | 2.445 | 2.53 | 2.4096 | 2.32 | 1.5518
| Std Dev. |  | 0.024946 | 0.052603 | 0.050426 | 0.065478 | 0.039867 | 0.067999 | 0.05164 | 0.024014 | 0.054772 | 0.12661 | 0.018348 | 0.13784 | 0.016342 | 0.029665 | 0.031644 | 0.041673 | 0.10627
| Rel.Std.Dev. |  | 1.03% | 2.103% | 1.983% | 2.515% | 1.647% | 2.799% | 2.094% | 0.9716% | 2.236% | 5.452% | 0.7364% | 5.626% | 0.668% | 1.168% | 1.321% | 1.787% | 6.667%
| PDM3 |  | -1.561% | 1.658% | 3.338% | 5.81% | -1.613% | -1.268% | 0.2357% | 0.4389% | -0.4415% | -5.638% | 1.252% | -0.4415% | -0.5905% | 3.216% | -2.647% | -5.25% | -35.22%
| Z-Score (Absolute) |  | 0.45 | 0.83 | 1.51 | 2.49 | 0.46 | 0.33 | 0.27 | 0.35 | 0 | 2.07 | 0.67 | 0 | 0.06 | 1.46 | 0.88 | 1.92 | 13.85
| Analytical results for La in OREAS 926b (Certified Value 35.7 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | [8 empty] | 1
2
| 1 | 1 | 36.7 | 37 | 34.8 | 29.5543 | 36.5 | 34.3 | 36 | 34.5 | 36.724 | 34.9 | [8 empty] | 1
| 1 | 2 | 34.5 | 37 | 34.5 | 29.7165 | 36.5 | 31.9 | 38 | 35.2 | 37.922 | 35.2 | [8 empty] | #N/A
| 1 | 3 | 34 | 38 | 35.3 | 29.9878 | 35.5 | 32.2 | 37 | 35.2 | 38.285 | 34.6 | [8 empty] | 16
| 1 | 4 | 34.7 | 36 | 35.1 | 29.7879 | 37 | 32.2 | 38 | 36.9 | 37.975 | 34.3 | [8 empty] | 35.6649
| 1 | 5 | 35.3 | 36 | 34.8 | 29.6096 | 35.5 | 30.9 | 37 | 34.8 | 38.128 | 35.6 | [8 empty] | 47
| 1 | 6 | 35.7 | 38 | 35.1 | 29.8606 | 35.5 | 32.8 | 38 | 34.9 | 38.769 | 34.7
| Mean |  | 35.15 | 37 | 34.9333 | 29.7528 | 36.0833 | 32.3833 | 37.3333 | 35.25 | 37.9672 | 34.8833
| Median |  | 35 | 37 | 34.95 | 29.7522 | 36 | 32.2 | 37.5 | 35.05 | 38.0515 | 34.8
| Std Dev. |  | 0.96695 | 0.89443 | 0.28752 | 0.16081 | 0.66458 | 1.1268 | 0.8165 | 0.85029 | 0.68079 | 0.46224
| Rel.Std.Dev. |  | 2.751% | 2.417% | 0.823% | 0.5405% | 1.842% | 3.48% | 2.187% | 2.412% | 1.793% | 1.325%
| PDM3 |  | -1.444% | 3.744% | -2.051% | -16.58% | 1.173% | -9.201% | 4.678% | -1.163% | 6.455% | -2.191%
| Z-Score (Absolute) |  | 0.03 | 0.9 | 0.13 | 2.74 | 0.44 | 1.42 | 1.07 | 0.03 | 1.39 | 0.16
| Analytical results for Li in OREAS 926b (Certified Value 24.4 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 30 | 30 | 30 | 20 | 30 | 24 | 25 | 24 | 22 | 20 | 20 | 22 | 19 | 20 |  |  |  |  | 1
| 1 | 2 | 30 | 30 | 30 | 20 | 30 | 21 | 25 | 25 | 22 | 20 | 20 | 22 | 18 | 21 |  |  |  |  | #N/A
| 1 | 3 | 30 | 30 | 30 | 20 | 30 | 21 | 23 | 26 | 23 | 40 | 20 | 22 | 20 | 20 |  |  |  |  | 16
| 1 | 4 | 30 | 30 | 30 | 20 | 30 | 22 | 23 | 24 | 23 | 30 | 20 | 22 | 18 | 21 |  |  |  |  | 24.3571
| 1 | 5 | 30 | 30 | 30 | 20 | 30 | 22 | 24 | 26 | 23 | 30 | 20 | 22 | 17 | 21 |  |  |  |  | 48
| 1 | 6 | 30 | 30 | 30 | 20 | 30 | 23 | 22 | 27 | 22 | 30 | 20 | 22 | 20 | 21
| Mean |  | 30 | 30 | 30 | 20 | 30 | 22.1667 | 23.6667 | 25.3333 | 22.5 | 28.3333 | 20 | 22 | 18.6667 | 20.6667
| Median |  | 30 | 30 | 30 | 20 | 30 | 22 | 23.5 | 25.5 | 22.5 | 30 | 20 | 22 | 18.5 | 21
| Std Dev. |  | 0 | 0 | 0 | 0 | 0 | 1.169 | 1.2111 | 1.2111 | 0.54772 | 7.5277 | 0 | 0 | 1.2111 | 0.5164
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 0% | 5.274% | 5.117% | 4.781% | 2.434% | 26.57% | 0% | 0% | 6.488% | 2.499%
| PDM3 |  | 23.17% | 23.17% | 23.17% | -17.89% | 23.17% | -8.993% | -2.835% | 4.008% | -7.625% | 16.32% | -17.89% | -9.677% | -23.36% | -15.15%
| Z-Score (Absolute) |  | 1.51 | 1.51 | 1.51 | 0.67 | 1.51 | 0.2 | 0.13 | 0.49 | 0.13 | 1.15 | 0.67 | 0.24 | 0.97 | 0.53
| Analytical results for Lu in OREAS 926b (Certified Value 0.35 ppm) | [19 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | < 2 | 0.42 | 0.4 | 0.35 | 0.37 | 0.35 | 0.347 | 0.34 | [10 empty] | 1
| 1 | 2 | < 2 | 0.43 | 0.4 | 0.32 | 0.37 | 0.33 | 0.347 | 0.36 | [10 empty] | 24
| 1 | 3 | < 2 | 0.43 | 0.4 | 0.32 | 0.35 | 0.35 | 0.351 | 0.35 | [10 empty] | 16
| 1 | 4 | < 2 | 0.41 | 0.8 | 0.32 | 0.4 | 0.34 | 0.342 | 0.31 | [10 empty] | 0.3452
| 1 | 5 | < 2 | 0.44 | 0.4 | 0.29 | 0.39 | 0.32 | 0.348 | 0.36 | [10 empty] | 49
| 1 | 6 | < 2 | 0.45 | 0.4 | 0.32 | 0.4 | 0.32 | 0.341 | 0.35
| Mean |  |  | 0.43 | 0.46667 | 0.32 | 0.38 | 0.335 | 0.346 | 0.345
| Median |  |  | 0.43 | 0.4 | 0.32 | 0.38 | 0.335 | 0.347 | 0.35
| Std Dev. |  |  | 0.014142 | 0.1633 | 0.018974 | 0.02 | 0.013784 | 0.0037947 | 0.018708
| Rel.Std.Dev. |  |  | 3.289% | 34.99% | 5.929% | 5.263% | 4.115% | 1.097% | 5.423%
| PDM3 |  |  | 24.57% | 35.19% | -7.3% | 10.08% | -2.955% | 0.2317% | -0.05794%
| Z-Score (Absolute) |  | 16.96 | 2.18 | NA | 0.67 | 0.88 | 0.29 | 0 | 0.03
Results from laboratory 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 926b (Certified Value 0.332 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | 3
3
| 1 | 1 | 0.32544 | 0.326 | 0.32 | 0.308 | 0.338 | 0.314 | 0.33 | 0.36 | 0.34 | 0.3383 | 0.34 | 0.34736 | 0.31 | 0.33 | 0.28 | 0.3227 | 0.35 | 0.3247 | 1
| 1 | 2 | 0.32908 | 0.32 | 0.32 | 0.32 | 0.332 | 0.314 | 0.33 | 0.36 | 0.34 | 0.3458 | 0.34 | 0.34297 | 0.33 | 0.33 | 0.26 | 0.3174 | 0.35 | 0.33373 | #N/A
| 1 | 3 | 0.32337 | 0.338 | 0.32 | 0.32 | 0.332 | 0.302 | 0.33 | 0.35 | 0.34 | 0.3659 | 0.35 | 0.34805 | 0.32 | 0.34 | 0.3 | 0.3169 | 0.34 | 0.3259 | 16
| 1 | 4 | 0.32809 | 0.326 | 0.338 | 0.32 | 0.332 | 0.32 | 0.33 | 0.35 | 0.36 | 0.3271 | 0.34 | 0.34701 | 0.33 | 0.34 | 0.28 | 0.3179 | 0.34 | 0.32892 | 0.33164
| 1 | 5 | 0.32786 | 0.32 | 0.326 | 0.308 | 0.338 | 0.314 | 0.33 | 0.35 | 0.35 | 0.3308 | 0.34 | 0.34646 | 0.31 | 0.34 | 0.27 | 0.316 | 0.34 | 0.30904 | 50
| 1 | 6 | 0.32871 | 0.326 | 0.326 | 0.302 | 0.338 | 0.308 | 0.33 | 0.35 | 0.35 | 0.3459 | 0.35 | 0.34394 | 0.33 | 0.31 | 0.3 | 0.3218 | 0.34 | 0.3247
| Mean |  | 0.32709 | 0.326 | 0.325 | 0.313 | 0.335 | 0.312 | 0.33 | 0.35333 | 0.34667 | 0.3423 | 0.34333 | 0.34597 | 0.32167 | 0.33167 | 0.28167 | 0.31878 | 0.34333 | 0.3245
| Median |  | 0.32797 | 0.326 | 0.323 | 0.314 | 0.335 | 0.314 | 0.33 | 0.35 | 0.345 | 0.34205 | 0.34 | 0.34674 | 0.325 | 0.335 | 0.28 | 0.31765 | 0.34 | 0.3253
| Std Dev. |  | 0.0022254 | 0.0065727 | 0.0070143 | 0.007975 | 0.0032863 | 0.0061968 | 0 | 0.005164 | 0.008165 | 0.013867 | 0.005164 | 0.0020359 | 0.0098319 | 0.01169 | 0.016021 | 0.0027723 | 0.005164 | 0.0083182
| Rel.Std.Dev. |  | 0.6804% | 2.016% | 2.158% | 2.548% | 0.981% | 1.986% | 0% | 1.462% | 2.355% | 4.051% | 1.504% | 0.5885% | 3.057% | 3.525% | 5.688% | 0.8697% | 1.504% | 2.563%
| PDM3 |  | -1.37% | -1.699% | -2.001% | -5.619% | 1.014% | -5.921% | -0.4934% | 6.542% | 4.532% | 3.216% | 3.527% | 4.321% | -3.006% | 0.009201% | -15.07% | -3.876% | 3.527% | -2.152%
| Z-Score (Absolute) |  | 0.08 | 0.17 | 0.22 | 0.92 | 0.35 | 0.97 | 0.08 | 1.42 | 1.04 | 0.79 | 0.85 | 1 | 0.39 | 0.18 | 2.69 | 0.56 | 0.85 | 0.23
| Analytical results for Mn in OREAS 926b (Certified Value 0.007 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | 3
3
| 1 | 1 | < 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0083 | 0.0073 | < 0.01 | 0.0083 | 0.007 | 0.012267 | 0.007 | < 0.01 | 0.012 | < 0.01 | 0.0072 | 0.0067933 | 1
| 1 | 2 | < 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0072 | < 0.01 | 0.0079 | 0.007 | 0.010566 | 0.006 | < 0.01 | 0.0119 | < 0.01 | 0.0069 | 0.0069746 | #N/A
| 1 | 3 | < 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0078 | 0.0071 | < 0.01 | 0.0081 | 0.007 | 0.0084172 | 0.006 | < 0.01 | 0.0118 | < 0.01 | 0.0069 | 0.0071269 | 16
| 1 | 4 | < 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0072 | < 0.01 | 0.0074 | 0.006 | 0.0092829 | 0.007 | < 0.01 | 0.0137 | < 0.01 | 0.0068 | 0.0072056 | 0.0074128
| 1 | 5 | < 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0079 | 0.0071 | < 0.01 | 0.0072 | 0.006 | 0.010644 | 0.007 | < 0.01 | 0.0126 | < 0.01 | 0.0068 | 0.0070366 | 51
| 1 | 6 | < 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0081 | 0.007 | < 0.01 | 0.0094 | 0.007 | 0.0099704 | 0.007 | < 0.01 | 0.0122 | < 0.01 | 0.0068 | 0.007527
| Mean |  |  | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0080167 | 0.00715 |  | 0.00805 | 0.0066667 | 0.010191 | 0.0066667 |  | 0.012367 |  | 0.0069 | 0.0071107
| Median |  |  | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.00715 |  | 0.008 | 0.007 | 0.010268 | 0.007 |  | 0.0121 |  | 0.00685 | 0.0070818
| Std Dev. |  |  | 0 | 0 | 0 | 0 | 0 | 0.00017224 | 0.00010488 |  | 0.00078166 | 0.0005164 | 0.0013172 | 0.0005164 |  | 0.00071181 |  | 0.00015492 | 0.00024791
| Rel.Std.Dev. |  |  | 0% | 0% | 0% | 0% | 0% | 2.149% | 1.467% |  | 9.71% | 7.746% | 12.93% | 7.746% |  | 5.756% |  | 2.245% | 3.486%
| PDM3 |  |  | 7.922% | 7.922% | 7.922% | 7.922% | 7.922% | 8.146% | -3.545% |  | 8.596% | -10.07% | 37.48% | -10.07% |  | 66.83% |  | -6.918% | -4.076%
| Z-Score (Absolute) |  | 2.82 | 0.34 | 0.34 | 0.34 | 0.34 | 0.34 | 0.66 | 0.34 | 2.82 | 0.69 | 0.9 | 3.16 | 0.9 | 2.82 | 5.66 | 2.82 | 0.63 | 0.39
| Analytical results for Mo in OREAS 926b (Certified Value 5.02 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 15 | 17 | 18 | 20 | 21 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 8 | 5 | 4 | 5 | 5 | 6 | 5 | 5 | 4.7583 | [9 empty] | 1
| 1 | 2 | 6 | 5 | 5 | 5 | 5 | 7 | 5 | 5 | 4.8512 | [9 empty] | 1
| 1 | 3 | 6 | 5 | 5 | 5 | < 5 | 6 | 5 | 5 | 4.8909 | [9 empty] | 16
| 1 | 4 | 6 | 5 | 5 | 5 | 5 | 5 | 6 | < 5 | 4.7658 | [9 empty] | 5.0157
| 1 | 5 | 5 | 6 | 5 | 5 | 5 | 5 | 6 | < 5 | 4.717 | [9 empty] | 52
| 1 | 6 | 7 | 5 | 5 | 5 | 5 | 5 | 5 | < 5 | 4.6765
| Mean |  | 6.3333 | 5.1667 | 4.8333 | 5 | 5 | 5.6667 | 5.3333 | 5 | 4.7766
| Median |  | 6 | 5 | 5 | 5 | 5 | 5.5 | 5 | 5 | 4.762
| Std Dev. |  | 1.0328 | 0.40825 | 0.40825 | 0 | 0 | 0.8165 | 0.5164 | 0 | 0.080831
| Rel.Std.Dev. |  | 16.31% | 7.902% | 8.447% | 0% | 0% | 14.41% | 9.682% | 0% | 1.692%
| PDM3 |  | 26.27% | 3.01% | -3.636% | -0.3132% | -0.3132% | 12.98% | 6.333% | -0.3132% | -4.767%
| Z-Score (Absolute) |  | 2.7 | 0.34 | 0.34 | 0 | 0.84 | 1.35 | 0.67 | 2.53 | 0.45
| Analytical results for Na in OREAS 926b (Indicative Value 0.084 wt.%) | [19 empty] | Indicative
Na | Round | Replicate | Lab | [17 empty] | 1
| No | No | 16 | [17 empty] | wt.%
|  |  | BF*OES/MS | [17 empty] | 3
3
| 1 | 1 | 0.089 | [17 empty] | 1
| 1 | 2 | 0.082 | [17 empty] | 14
| 1 | 3 | 0.089 | [17 empty] | 16
| 1 | 4 | 0.082 | [17 empty] | 0.084075
| 1 | 5 | 0.082 | [17 empty] | 20
| 1 | 6 | 0.082
| Mean |  | 0.084333
| Median |  | 0.082
| Std Dev. |  | 0.0036148
| Rel.Std.Dev. |  | 4.286%
| PDM3 |  | 0.3071%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 926b (Certified Value 10.2 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 12 | 15 | 16 | 17 | 18 | 20 | 21 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [10 empty] | 1
1
| 1 | 1 | 10 | 8 | 10 | 10.4 | 12.6 | < 10 | 12 | 11.1355 | [10 empty] | 1
| 1 | 2 | 10.1 | 8 | 10 | 9.5 | 12.17 | < 10 | 14 | 12.0233 | [10 empty] | 3
| 1 | 3 | 10 | 9 | 10 | 9.7 | 12.56 | 10 | 13 | 10.178 | [10 empty] | 16
| 1 | 4 | 10.1 | 30 | 10 | 9.3 | 13.6 | < 10 | 14 | 11.7381 | [10 empty] | 10.218
| 1 | 5 | 10.1 | 9 | 10 | 8.9 | 11.76 | < 10 | 13 | 11.5456 | [10 empty] | 53
| 1 | 6 | 10.6 | 8 | 10 | 9.3 | 11.76 | < 10 | 13 | 9.6864
| Mean |  | 10.15 | 12 | 10 | 9.5167 | 12.4083 | 10 | 13.1667 | 11.0511
| Median |  | 10.1 | 8.5 | 10 | 9.4 | 12.365 | 10 | 13 | 11.3406
| Std Dev. |  | 0.22583 | 8.8318 | 0 | 0.50761 | 0.68953 |  | 0.75277 | 0.92667
| Rel.Std.Dev. |  | 2.225% | 73.6% | 0% | 5.334% | 5.557% |  | 5.717% | 8.385%
| PDM3 |  | -0.6657% | 17.44% | -2.134% | -6.864% | 21.44% | -2.134% | 28.86% | 8.153%
| Z-Score (Absolute) |  | 0.24 | 0.76 | 0.33 | 0.59 | 0.98 | 2.59 | 1.39 | 0.24
| Analytical results for Nd in OREAS 926b (Certified Value 29.2 ppm) | [19 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [8 empty] | 1
2
| 1 | 1 | 30.8 | 27 | 28.9 | 28.5466 | 32 | 29.4 | 29.5 | 29.4 | 25.655 | 29 | [8 empty] | 1
| 1 | 2 | 29 | 28 | 29.4 | 28.202 | 32.5 | 27.4 | 30.9 | 29.4 | 26.511 | 31 | [8 empty] | 4
| 1 | 3 | 28.4 | 29 | 29.8 | 28.1447 | 31 | 27.6 | 31.7 | 30.9 | 26.762 | 31 | [8 empty] | 16
| 1 | 4 | 29.1 | 27 | 28.8 | 28.3467 | 32 | 27.7 | 31.8 | 30.8 | 26.065 | 30 | [8 empty] | 29.2436
| 1 | 5 | 31.4 | 27 | 30 | 28.1212 | 31 | 26.1 | 31.4 | 29.8 | 26.418 | 31 | [8 empty] | 54
| 1 | 6 | 31 | 30 | 29.9 | 28.5388 | 31.5 | 28.4 | 28.6 | 30 | 26.003 | 30
| Mean |  | 29.95 | 28 | 29.4667 | 28.3167 | 31.6667 | 27.7667 | 30.65 | 30.05 | 26.2357 | 30.3333
| Median |  | 29.95 | 27.5 | 29.6 | 28.2744 | 31.75 | 27.65 | 31.15 | 29.9 | 26.2415 | 30.5
| Std Dev. |  | 1.2613 | 1.2649 | 0.52026 | 0.19184 | 0.60553 | 1.0967 | 1.3096 | 0.66257 | 0.40163 | 0.8165
| Rel.Std.Dev. |  | 4.212% | 4.518% | 1.766% | 0.6775% | 1.912% | 3.95% | 4.273% | 2.205% | 1.531% | 2.692%
| PDM3 |  | 2.416% | -4.252% | 0.7629% | -3.17% | 8.286% | -5.05% | 4.809% | 2.758% | -10.29% | 3.727%
| Z-Score (Absolute) |  | 0.14 | 0.99 | 0.14 | 0.8 | 1.13 | 1.12 | 0.54 | 0.2 | 2.01 | 0.36
| Analytical results for Ni in OREAS 926b (Indicative Value 17 ppm) | [19 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS |  |  | 1
1
| 1 | 1 | < 50 | < 50 | 80 | < 50 | < 50 | 20 | 20 | < 50 | 9 | < 20 | < 50 | < 50 | 21 | < 10 | < 10 | 21.7138 |  |  | 1
| 1 | 2 | < 50 | < 50 | 70 | < 50 | < 50 | 10 | 27 | < 50 | 20 | < 20 | < 50 | < 50 | 21 | < 10 | < 10 | 19.8073 |  |  | 5
| 1 | 3 | < 50 | < 50 | 80 | < 50 | < 50 | 50 | < 20 | < 50 | 12 | < 20 | < 50 | < 50 | 22 | 12 | 16 | 22.9686 |  |  | 16
| 1 | 4 | < 50 | < 50 | < 50 | < 50 | < 50 | 30 | 48 | < 50 | 13 | < 20 | < 50 | < 50 | 20 | 20 | < 10 | 23.4534 |  |  | 16.9911
| 1 | 5 | < 50 | < 50 | 90 | < 50 | < 50 | 10 | 36 | < 50 | 16 | < 20 | < 50 | < 50 | 23 | 10 | < 10 | 21.2681 |  |  | 17
| 1 | 6 | < 50 | < 50 | 180 | < 50 | < 50 | 20 | 31 | < 50 | 11 | < 20 | < 50 | < 50 | 21 | 14 | < 10 | 19.4144
| Mean |  |  |  | 100 |  |  | 23.3333 | 32.4 |  | 13.5 |  |  |  | 21.3333 | 14 | 16 | 21.4376
| Median |  |  |  | 80 |  |  | 20 | 31 |  | 12.5 |  |  |  | 21 | 13 | 16 | 21.491
| Std Dev. |  |  |  | 45.2769 |  |  | 15.0555 | 10.5024 |  | 3.937 |  |  |  | 1.0328 | 4.3205 |  | 1.6287
| Rel.Std.Dev. |  |  |  | 45.28% |  |  | 64.52% | 32.41% |  | 29.16% |  |  |  | 4.841% | 30.86% |  | 7.597%
| PDM3 |  |  |  | 488.5% |  |  | 37.33% | 90.69% |  | -20.55% |  |  |  | 25.56% | -17.6% | -5.833% | 26.17%
| Z-Score (Absolute) |  | 0 | 0 | 16.12 | 0 | 0 | 0.43 | 0.95 | 0 | 2.97 | 3.87 | 0 | 0 | 0.95 | 3.61 | 4.69 | 0.92
| Analytical results for P in OREAS 926b (Certified Value 0.029 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | [10 empty] | wt.%
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [10 empty] | 3
3
| 1 | 1 | 0.025068 | 0.0358 | 0.03 | 0.01 | 0.026 | 0.03 | 0.0268 | 0.03 | [10 empty] | 1
| 1 | 2 | 0.025101 | 0.0289 | 0.03 | 0.01 | 0.022 | 0.03 | 0.0281 | 0.03 | [10 empty] | #N/A
| 1 | 3 | 0.02553 | 0.0295 | 0.03 | 0.01 | 0.026 | 0.03 | 0.0308 | 0.02 | [10 empty] | 16
| 1 | 4 | 0.025356 | 0.0377 | 0.03 | 0.02 | 0.026 | 0.03 | 0.0244 | 0.03 | [10 empty] | 0.028574
| 1 | 5 | 0.024348 | 0.034 | 0.03 | 0.01 | 0.026 | 0.03 | 0.0344 | 0.02 | [10 empty] | 55
| 1 | 6 | 0.02549 | 0.0327 | 0.03 | 0.01 | 0.026 | 0.03 | 0.0334 | 0.03
| Mean |  | 0.025149 | 0.0331 | 0.03 | 0.011667 | 0.025333 | 0.03 | 0.02965 | 0.026667
| Median |  | 0.025229 | 0.03335 | 0.03 | 0.01 | 0.026 | 0.03 | 0.02945 | 0.03
| Std Dev. |  | 0.00043698 | 0.0034647 | 0 | 0.0040825 | 0.001633 | 0 | 0.0038996 | 0.005164
| Rel.Std.Dev. |  | 1.738% | 10.47% | 0% | 34.99% | 6.446% | 0% | 13.15% | 19.36%
| PDM3 |  | -11.99% | 15.84% | 4.989% | -59.17% | -11.34% | 4.989% | 3.764% | -6.676%
| Z-Score (Absolute) |  | 0.89 | 1.47 | 0.55 | 4.89 | 0.8 | 0.55 | 0.44 | 0.44
| Analytical results for Pb in OREAS 926b (Certified Value 51 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | < 100 | 100 | 100 | 100 | 100 | 100 | 46.7 | 65 | 300 | 50 | 50 | < 300 | 55 | 80 | 63 | 49.6577 |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | 100 | 100 | 49.4 | 58 | < 100 | 50 | 50 | < 300 | 58 | 99 | 50 | 45.3246 |  |  | 6
| 1 | 3 | < 100 | < 100 | 100 | < 100 | 100 | 100 | 45.9 | 66 | < 100 | 50 | 40 | < 300 | 54 | 59 | 48 | 46.164 |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | 100 | 46.7 | 51 | < 100 | 50 | 50 | < 300 | 54 | 62 | 49 | 45.5927 |  |  | 51.0648
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | 100 | 45 | 52 | < 100 | 50 | 40 | < 300 | 55 | 55 | 50 | 45.1373 |  |  | 56
| 1 | 6 | < 100 | < 100 | 100 | < 100 | 100 | < 100 | 46.5 | 52 | < 100 | 50 | 50 | < 300 | 56 | 53 | 48 | 45.5329
| Mean |  |  | 100 | 100 | 100 | 100 | 100 | 46.7 | 57.3333 | 300 | 50 | 46.6667 |  | 55.3333 | 68 | 51.3333 | 46.2349
| Median |  |  | 100 | 100 | 100 | 100 | 100 | 46.6 | 55 | 300 | 50 | 50 |  | 55 | 60.5 | 49.5 | 45.5628
| Std Dev. |  |  |  | 0 |  | 0 | 0 | 1.4738 | 6.802 |  | 0 | 5.164 |  | 1.5055 | 17.9778 | 5.785 | 1.7123
| Rel.Std.Dev. |  |  |  | 0% |  | 0% | 0% | 3.156% | 11.86% |  | 0% | 11.07% |  | 2.721% | 26.44% | 11.27% | 3.703%
| PDM3 |  |  | 95.83% | 95.83% | 95.83% | 95.83% | 95.83% | -8.548% | 12.28% | 487.5% | -2.085% | -8.613% |  | 8.359% | 33.16% | 0.5259% | -9.458%
| Z-Score (Absolute) |  | 0.19 | NA | NA | NA | NA | NA | 0.67 | 0.87 | 5.83 | 0.19 | 0.67 | 14.26 | 0.58 | 2.41 | 0 | 0.74
Results from laboratories 3, 4, 5, 6 and 7 were removed due to their 100 ppm reading resolution.
| Analytical results for Pr in OREAS 926b (Certified Value 7.85 ppm) | [19 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 8.7 | 7 | 7.74 | 6.5311 | 8.4 | 7.92 | 8.5 | 8.05 | 6.913 | 7.97 | [8 empty] | 1
| 1 | 2 | 8.1 | 7 | 7.87 | 6.6468 | 8.6 | 7.31 | 8.6 | 8.3 | 7.141 | 8.08 | [8 empty] | 7
| 1 | 3 | 8.3 | 8 | 7.93 | 6.7078 | 8.2 | 7.34 | 8.6 | 8.2 | 7.194 | 7.76 | [8 empty] | 16
| 1 | 4 | 8.4 | 7 | 8.09 | 6.5634 | 8.8 | 7.45 | 8.5 | 8.39 | 7.056 | 7.91 | [8 empty] | 7.845
| 1 | 5 | 8.2 | 7 | 7.83 | 6.8813 | 8.4 | 7.13 | 8.3 | 8.23 | 7.121 | 8.05 | [8 empty] | 57
| 1 | 6 | 8.3 | 8 | 7.78 | 6.8157 | 8.4 | 7.59 | 8.7 | 8.25 | 6.971 | 7.92
| Mean |  | 8.3333 | 7.3333 | 7.8733 | 6.691 | 8.4667 | 7.4567 | 8.5333 | 8.2367 | 7.066 | 7.9483
| Median |  | 8.3 | 7 | 7.85 | 6.6773 | 8.4 | 7.395 | 8.55 | 8.24 | 7.0885 | 7.945
| Std Dev. |  | 0.20656 | 0.5164 | 0.12533 | 0.13846 | 0.20656 | 0.27362 | 0.13663 | 0.11308 | 0.1073 | 0.11479
| Rel.Std.Dev. |  | 2.479% | 7.042% | 1.592% | 2.069% | 2.44% | 3.669% | 1.601% | 1.373% | 1.519% | 1.444%
| PDM3 |  | 6.224% | -6.523% | 0.3607% | -14.71% | 7.924% | -4.951% | 8.774% | 4.992% | -9.93% | 1.317%
| Z-Score (Absolute) |  | 0.53 | NA | 0.1 | 1.72 | 0.71 | 0.67 | 0.8 | 0.4 | 1.21 | 0
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 926b (Certified Value 141 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | 136 | 144.1 | 140 | 119.4558 | 151 | 150.5 | 148 | 138 | 136 | 131.7586 | [8 empty] | 1
| 1 | 2 | 138 | 141.5 | 142 | 119.644 | 152 | 140.2 | 150 | 140 | 139 | 132.7723 | [8 empty] | 8
| 1 | 3 | 133 | 142.6 | 141 | 119.3022 | 145 | 142.9 | 151 | 143 | 133 | 135.2593 | [8 empty] | 16
| 1 | 4 | 137 | 142 | 141 | 119.7733 | 139 | 138.3 | 145 | 143 | 136 | 134.8963 | [8 empty] | 140.5121
| 1 | 5 | 131 | 143.2 | 142 | 119.5964 | 145 | 134 | 151 | 141 | 137 | 134.3091 | [8 empty] | 58
| 1 | 6 | 133 | 142.3 | 143 | 120.0687 | 142 | 141.5 | 142 | 141 | 136 | 134.5564
| Mean |  | 134.6667 | 142.6167 | 141.5 | 119.6401 | 145.6667 | 141.2333 | 147.8333 | 141 | 136.1667 | 133.9254
| Median |  | 134.5 | 142.45 | 141.5 | 119.6202 | 145 | 140.85 | 149 | 141 | 136 | 134.4328
| Std Dev. |  | 2.7325 | 0.92394 | 1.0488 | 0.26514 | 5.0465 | 5.4902 | 3.656 | 1.8974 | 1.9408 | 1.3632
| Rel.Std.Dev. |  | 2.029% | 0.6478% | 0.7412% | 0.2216% | 3.464% | 3.887% | 2.473% | 1.346% | 1.425% | 1.018%
| PDM3 |  | -4.16% | 1.498% | 0.7031% | -14.85% | 3.668% | 0.5133% | 5.21% | 0.3472% | -3.093% | -4.688%
| Z-Score (Absolute) |  | 0.92 | 0.21 | 0.05 | 3.05 | 0.65 | 0.02 | 0.95 | 0.02 | 0.7 | 1.02
| Analytical results for Re in OREAS 926b (Indicative Value < 0.1 ppm) | [19 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 10 | 13 | 15 | 17 | [14 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | [14 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | < 1 | [14 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 1 | < 1 | [14 empty] | 12
| 1 | 3 | < 0.1 | < 0.1 | < 1 | < 1 | [14 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | < 1 | [14 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | < 1 | [14 empty] | 18
| 1 | 6 | < 0.1 | < 0.1 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for S in OREAS 926b (Certified Value 0.981 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 15 | 16 | 17 | 20 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.98045 | 1.01 | 1.03 | 1.04 | 1 | 1 | 0.95 | 0.98 | 0.9007 | 1 | 0.92 | 0.91 | 1.06 | 0.92 |  |  |  |  | 1
| 1 | 2 | 0.98994 | 1.02 | 1.06 | 1.08 | 1 | 1 | 0.93 | 0.97 | 0.8767 | 0.98 | 0.97 | 0.93 | 1.07 | 0.93 |  |  |  |  | #N/A
| 1 | 3 | 0.97595 | 1.02 | 1.06 | 1.06 | 1 | 0.94 | 0.94 | 0.96 | 0.9401 | 0.99 | 0.92 | 0.94 | 1.06 | 0.9 |  |  |  |  | 16
| 1 | 4 | 0.98365 | 0.99 | 1.07 | 1.1 | 1.01 | 0.99 | 0.94 | 0.97 | 0.87 | 0.97 | 0.93 | 0.93 | 1.06 | 0.91 |  |  |  |  | 0.98054
| 1 | 5 | 0.9845 | 0.99 | 1.04 | 1.08 | 1.02 | 0.98 | 0.96 | 0.98 | 0.918 | 0.97 | 0.9 | 0.94 | 1.05 | 0.92 |  |  |  |  | 59
| 1 | 6 | 0.99232 | 1.01 | 1.04 | 1.02 | 1.01 | 0.97 | 0.94 | 0.96 | 0.9433 | 0.97 | 0.94 | 0.91 | 1.06 | 0.93
| Mean |  | 0.98447 | 1.0067 | 1.05 | 1.0633 | 1.0067 | 0.98 | 0.94333 | 0.97 | 0.90813 | 0.98 | 0.93 | 0.92667 | 1.06 | 0.91833
| Median |  | 0.98408 | 1.01 | 1.05 | 1.07 | 1.005 | 0.985 | 0.94 | 0.97 | 0.90935 | 0.975 | 0.925 | 0.93 | 1.06 | 0.92
| Std Dev. |  | 0.0060176 | 0.013663 | 0.015492 | 0.029439 | 0.008165 | 0.022804 | 0.010328 | 0.0089443 | 0.031156 | 0.012649 | 0.023664 | 0.013663 | 0.0063246 | 0.01169
| Rel.Std.Dev. |  | 0.6113% | 1.357% | 1.475% | 2.769% | 0.8111% | 2.327% | 1.095% | 0.9221% | 3.431% | 1.291% | 2.545% | 1.474% | 0.5967% | 1.273%
| PDM3 |  | 0.4005% | 2.664% | 7.084% | 8.443% | 2.664% | -0.05539% | -3.795% | -1.075% | -7.385% | -0.05539% | -5.155% | -5.495% | 8.103% | -6.344%
| Z-Score (Absolute) |  | 0.07 | 0.41 | 1.09 | 1.3 | 0.41 | 0 | 0.57 | 0.16 | 1.12 | 0 | 0.78 | 0.83 | 1.24 | 0.96
| Analytical results for Sb in OREAS 926b (Certified Value 4.98 ppm) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 10 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 5 | 4.4 | 5.5 | 4.279 | 10 | < 5 | 5 | 6 | 5.7128 | [9 empty] | 1
| 1 | 2 | 5 | 4.6 | 5.2 | 4.0967 | 10 | < 5 | 5.1 | 5 | 5.6701 | [9 empty] | 9
| 1 | 3 | 5 | 4.7 | 5.3 | 4.316 | < 10 | < 5 | 5 | 5 | 5.722 | [9 empty] | 16
| 1 | 4 | 5 | 4.5 | 5.5 | 4.179 | < 10 | < 5 | 4.9 | 5 | 5.6664 | [9 empty] | 4.9767
| 1 | 5 | 5 | 4.4 | 5.3 | 4.1463 | < 10 | < 5 | 5 | 5 | 5.6515 | [9 empty] | 60
| 1 | 6 | 5 | 4.8 | 5.5 | 4.3925 | < 10 | < 5 | 4.7 | 5 | 6.0677
| Mean |  | 5 | 4.5667 | 5.3833 | 4.2349 | 10 |  | 4.95 | 5.1667 | 5.7484
| Median |  | 5 | 4.55 | 5.4 | 4.229 | 10 |  | 5 | 5 | 5.6915
| Std Dev. |  | 0 | 0.1633 | 0.13292 | 0.11263 | 0 |  | 0.13784 | 0.40825 | 0.15882
| Rel.Std.Dev. |  | 0% | 3.576% | 2.469% | 2.659% | 0% |  | 2.785% | 7.902% | 2.763%
| PDM3 |  | 0.4688% | -8.238% | 8.171% | -14.9% | 100.9% |  | -0.5359% | 3.818% | 15.51%
| Z-Score (Absolute) |  | NA | 0.23 | 0.73 | 0.61 | NA | 2.65 | 0.23 | NA | 1.16
| Analytical results for Sc in OREAS 926b (Indicative Value 9.03 ppm) | [19 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 10 | 8 | < 20 | < 10 | 10 | 15 | 9 | 9.486 | < 10 | [9 empty] | 1
| 1 | 2 | < 10 | 9 | < 20 | < 10 | 10 | 15 | 9 | 9.243 | < 10 | [9 empty] | 13
| 1 | 3 | < 10 | 9 | < 20 | < 10 | 10 | 15 | 8 | 9.356 | < 10 | [9 empty] | 16
| 1 | 4 | < 10 | 8 | < 20 | < 10 | 9 | 16 | 9 | 9.135 | < 10 | [9 empty] | 9.0298
| 1 | 5 | < 10 | 8 | < 20 | < 10 | 9 | 17 | 9 | 9.154 | < 10 | [9 empty] | 19
| 1 | 6 | < 10 | 9 | < 20 | < 10 | 9 | 16 | 9 | 9.341 | < 10
| Mean |  |  | 8.5 |  |  | 9.5 | 15.6667 | 8.8333 | 9.2858
| Median |  |  | 8.5 |  |  | 9.5 | 15.5 | 9 | 9.292
| Std Dev. |  |  | 0.54772 |  |  | 0.54772 | 0.8165 | 0.40825 | 0.13416
| Rel.Std.Dev. |  |  | 6.444% |  |  | 5.766% | 5.212% | 4.622% | 1.445%
| PDM3 |  |  | -5.867% |  |  | 5.207% | 73.5% | -2.176% | 2.836%
| Z-Score (Absolute) |  | 2.22 | 0.19 | 0.67 | 2.22 | 0.39 | 3.95 | 0 | 0.26 | 2.22
| Analytical results for Se in OREAS 926b (Indicative Value 10.3 ppm) | [19 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 08 | 10 | 13 | 20 | 21 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES/MS | [13 empty] | 1
1
| 1 | 1 | 15 | < 20 | < 20 | 13 | 4.9018 | [13 empty] | 1
| 1 | 2 | 16 | < 20 | < 20 | 12 | 5.0819 | [13 empty] | 14
| 1 | 3 | 14 | < 20 | < 20 | 12 | 5.2335 | [13 empty] | 16
| 1 | 4 | 13 | < 20 | < 20 | 10 | 4.8186 | [13 empty] | 10.3004
| 1 | 5 | 15 | < 20 | < 20 | 11 | 5.1294 | [13 empty] | 20
| 1 | 6 | 12 | < 20 | < 20 | 12 | 5.2424
| Mean |  | 14.1667 |  |  | 11.6667 | 5.0679
| Median |  | 14.5 |  |  | 12 | 5.1057
| Std Dev. |  | 1.472 |  |  | 1.0328 | 0.17414
| Rel.Std.Dev. |  | 10.39% |  |  | 8.853% | 3.436%
| PDM3 |  | 37.53% |  |  | 13.26% | -50.8%
| Z-Score (Absolute) |  | 1.69 | 0 | 0 | 0.67 | 2
| Analytical results for Si in OREAS 926b (Certified Value 35.24 wt.%) | [19 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 20 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS |  |  |  | 2
3
| 1 | 1 | 36.4485 | 34.69 | 35.81 | 35.11 | 35.2 | 35.48 | 35.5 | 35.6 | 32.3 | > 30 | 34.0807 | 34.7 | 34.611 | 35.6 | 34.8 |  |  |  | 1
| 1 | 2 | 36.6459 | 34.08 | 36.18 | 37.91 | 35.01 | 35.48 | 35.4 | 35.6 | 34.1 | > 30 | 34.1187 | 35 | 34.686 | 35.4 | 34.7 |  |  |  | #N/A
| 1 | 3 | 36.5237 | 35.71 | 36.6 | 38.99 | 35.39 | 33.28 | 35.9 | 35.2 | 34.3 | > 30 | 34.1349 | 35 | 34.611 | 34.6 | 33.9 |  |  |  | 16
| 1 | 4 | 36.5707 | 34.92 | 37.12 | 38.29 | 35.29 | 35.39 | 35.2 | 35.1 | 32.5 | > 30 | 34.096 | 34.7 | 34.555 | 34.7 | 34.2 |  |  |  | 35.241
| 1 | 5 | 36.5143 | 34.59 | 37.21 | 34.31 | 35.29 | 35.01 | 35.2 | 35.3 | 32.5 | > 30 | 34.2332 | 34 | 34.695 | 35 | 34.4 |  |  |  | 61
| 1 | 6 | 36.4955 | 35.01 | 36.7 | 33.61 | 35.48 | 34.78 | 35.5 | 34.9 | 32.4 | > 30 | 34.0608 | 35.1 | 34.588 | 35.5 | 34.7
| Mean |  | 36.5331 | 34.8333 | 36.6033 | 36.37 | 35.2767 | 34.9033 | 35.45 | 35.2833 | 33.0167 |  | 34.1207 | 34.75 | 34.6243 | 35.1333 | 34.45
| Median |  | 36.519 | 34.805 | 36.65 | 36.51 | 35.29 | 35.2 | 35.45 | 35.25 | 32.5 |  | 34.1074 | 34.85 | 34.611 | 35.2 | 34.55
| Std Dev. |  | 0.068045 | 0.53936 | 0.53906 | 2.2966 | 0.16219 | 0.84424 | 0.25884 | 0.27869 | 0.92177 |  | 0.061079 | 0.40373 | 0.05528 | 0.4274 | 0.35071
| Rel.Std.Dev. |  | 0.1863% | 1.548% | 1.473% | 6.314% | 0.4598% | 2.419% | 0.7302% | 0.7899% | 2.792% |  | 0.179% | 1.162% | 0.1597% | 1.216% | 1.018%
| PDM3 |  | 3.667% | -1.157% | 3.866% | 3.204% | 0.1013% | -0.958% | 0.5932% | 0.1202% | -6.312% |  | -3.179% | -1.393% | -1.75% | -0.3054% | -2.244%
| Z-Score (Absolute) |  | 2.47 | 0.3 | 2.59 | 2.21 | 0.42 | 0.19 | 0.7 | 0.43 | 3.27 | NA | 1.47 | 0.44 | 0.64 | 0.19 | 0.93
| Analytical results for Sm in OREAS 926b (Certified Value 5.51 ppm) | [19 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 5.5 | 5 | 4.7078 | 5 | 5.65 | 5.8 | 5.7 | 5.747 | 5.6 | [9 empty] | 1
| 1 | 2 | 5.4 | 5 | 4.805 | 5 | 5.02 | 6.1 | 5.8 | 5.984 | 5.7 | [9 empty] | 1
| 1 | 3 | 5.3 | 6 | 4.8909 | 5 | 5.25 | 6.2 | 6 | 5.951 | 5.7 | [9 empty] | 16
| 1 | 4 | 5 | 5 | 4.7384 | 5.5 | 5.24 | 5.9 | 6.1 | 6.009 | 5.9 | [9 empty] | 5.5091
| 1 | 5 | 5.6 | 5 | 4.5774 | 5 | 4.94 | 6.1 | 5.9 | 6.032 | 6.3 | [9 empty] | 62
| 1 | 6 | 5.6 | 6 | 4.8556 | 5 | 5.22 | 6.1 | 5.9 | 6 | 5.7
| Mean |  | 5.4 | 5.3333 | 4.7625 | 5.0833 | 5.22 | 6.0333 | 5.9 | 5.9538 | 5.8167
| Median |  | 5.45 | 5 | 4.7717 | 5 | 5.23 | 6.1 | 5.9 | 5.992 | 5.7
| Std Dev. |  | 0.22804 | 0.5164 | 0.11384 | 0.20412 | 0.24682 | 0.15055 | 0.14142 | 0.10486 | 0.25626
| Rel.Std.Dev. |  | 4.223% | 9.682% | 2.39% | 4.016% | 4.728% | 2.495% | 2.397% | 1.761% | 4.406%
| PDM3 |  | -1.981% | -3.191% | -13.55% | -7.729% | -5.248% | 9.515% | 7.095% | 8.072% | 5.582%
| Z-Score (Absolute) |  | 0.38 | NA | 1.55 | 0.96 | 0.71 | 0.78 | 0.54 | 0.63 | 0.38
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 926b (Certified Value 13.3 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [9 empty] | 1
1
| 1 | 1 | 11.8 | < 100 | < 50 | < 20 | < 10 | 13 | 16 | 12 | 13 | [9 empty] | 1
| 1 | 2 | 12.3 | < 100 | < 50 | < 20 | < 10 | 12 | 16 | 13 | 14 | [9 empty] | 2
| 1 | 3 | 11.7 | < 100 | < 50 | < 20 | < 10 | 13 | 16 | 13 | 13 | [9 empty] | 16
| 1 | 4 | 11.8 | < 100 | < 50 | < 20 | < 10 | 15 | 16 | 13 | 13 | [9 empty] | 13.3433
| 1 | 5 | 11.2 | < 100 | < 50 | < 20 | < 10 | 12 | 16 | 13 | 13 | [9 empty] | 63
| 1 | 6 | 12.5 | < 100 | < 50 | < 20 | < 10 | 13 | 16 | 12 | 13
| Mean |  | 11.8833 |  |  |  |  | 13 | 16 | 12.6667 | 13.1667
| Median |  | 11.8 |  |  |  |  | 13 | 16 | 13 | 13
| Std Dev. |  | 0.46224 |  |  |  |  | 1.0954 | 0 | 0.5164 | 0.40825
| Rel.Std.Dev. |  | 3.89% |  |  |  |  | 8.427% | 0% | 4.077% | 3.101%
| PDM3 |  | -10.94% |  |  |  |  | -2.573% | 19.91% | -5.071% | -1.324%
| Z-Score (Absolute) |  | 0.25 | 8.32 | 2.7 | 0.67 | 1.8 | 0 | 0.67 | 0.07 | 0.04
| Analytical results for Sr in OREAS 926b (Certified Value 45.6 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 08 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 42.61 | 50 | 46 | 50 | 51 | 40 | 46 | 46 | 46 | 34 | 42 | 50.4393 | [6 empty] | 1
| 1 | 2 | 44.04 | 49 | 46 | < 50 | 52 | 40 | 47 | 42.1 | 46 | 34 | 40 | 53.1751 | [6 empty] | 3
| 1 | 3 | 43.36 | 49 | 47 | 50 | 54 | 40 | 48 | 43 | 49 | 35 | 40 | 52.3732 | [6 empty] | 16
| 1 | 4 | 43.96 | 48 | 47 | 50 | 49 | 40 | 46 | 42.3 | 47 | 34 | 40 | 53.6157 | [6 empty] | 45.6259
| 1 | 5 | 42.92 | 50 | 48 | < 50 | 52 | 40 | 47 | 40.8 | 48 | 35 | 40 | 51.6126 | [6 empty] | 64
| 1 | 6 | 43.92 | 50 | 47 | 60 | 53 | 40 | 45 | 42.8 | 48 | 35 | 40 | 51.0405
| Mean |  | 43.4683 | 49.3333 | 46.8333 | 52.5 | 51.8333 | 40 | 46.5 | 42.8333 | 47.3333 | 34.5 | 40.3333 | 52.0427
| Median |  | 43.64 | 49.5 | 47 | 50 | 52 | 40 | 46.5 | 42.55 | 47.5 | 34.5 | 40 | 51.9929
| Std Dev. |  | 0.6036 | 0.8165 | 0.75277 | 5 | 1.7224 | 0 | 1.0488 | 1.7328 | 1.2111 | 0.54772 | 0.8165 | 1.235
| Rel.Std.Dev. |  | 1.389% | 1.655% | 1.607% | 9.524% | 3.323% | 0% | 2.256% | 4.045% | 2.559% | 1.588% | 2.024% | 2.373%
| PDM3 |  | -4.729% | 8.126% | 2.646% | 15.07% | 13.61% | -12.33% | 1.916% | -6.121% | 3.742% | -24.39% | -11.6% | 14.06%
| Z-Score (Absolute) |  | 0.31 | 0.88 | 0.37 | 0.33 | 1.38 | 1 | 0.31 | 0.43 | 0.47 | 2.11 | 0.94 | 1.42
| Analytical results for Ta in OREAS 926b (Indicative Value 1.06 ppm) | [19 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 1.4 | 1 | < 10 | 0.2 | 1.5 | 0.8 | < 1 | < 10 | 0.7 | 5 | [8 empty] | 1
| 1 | 2 | 0.7 | 0.7 | < 10 | 0.2 | 2 | 0.8 | < 1 | < 10 | 1.6 | 5.5114 | [8 empty] | 4
| 1 | 3 | 0.8 | 1 | < 10 | 0.2 | 1.5 | 0.7 | < 1 | < 10 | < 0.5 | 5.884 | [8 empty] | 16
| 1 | 4 | 0.8 | 1 | < 10 | 0.2 | 1.5 | 0.7 | < 1 | < 10 | < 0.5 | 5.8379 | [8 empty] | 1.0556
| 1 | 5 | 1.2 | 0.9 | < 10 | 0.2 | 1 | 0.7 | 1 | < 10 | < 0.5 | 5.0849 | [8 empty] | 17
| 1 | 6 | 1.4 | 0.8 | < 10 | 0.2 | 1.5 | 0.7 | < 1 | < 10 | < 0.5 | 5.4811
| Mean |  | 1.05 | 0.9 |  | 0.2 | 1.5 | 0.73333 | 1 |  | 1.15 | 5.4665
| Median |  | 1 | 0.95 |  | 0.2 | 1.5 | 0.7 | 1 |  | 1.15 | 5.4963
| Std Dev. |  | 0.32094 | 0.12649 |  | 3.0405e-17 | 0.31623 | 0.05164 |  |  | 0.6364 | 0.36813
| Rel.Std.Dev. |  | 30.57% | 14.05% |  | 1.52e-14% | 21.08% | 7.042% |  |  | 55.34% | 6.734%
| PDM3 |  | -0.5263% | -14.74% |  | -81.05% | 42.11% | -30.53% | -5.263% |  | 8.947% | 417.9%
| Z-Score (Absolute) |  | 0.11 | 0.11 | 5.71 | 1.1 | 0.75 | 0.34 | 0.56 | 5.71 | 0.6 | 6.38
| Analytical results for Tb in OREAS 926b (Certified Value 0.69 ppm) | [19 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.7 | < 3 | 0.65 | 0.6 | 0.72 | 0.72 | 0.7 | 0.68 | 0.67 | [9 empty] | 1
| 1 | 2 | 0.7 | < 3 | 0.66 | 0.6 | 0.66 | 0.73 | 0.73 | 0.706 | 0.71 | [9 empty] | 21
| 1 | 3 | 0.7 | < 3 | 0.65 | 0.6 | 0.66 | 0.76 | 0.71 | 0.713 | 0.69 | [9 empty] | 16
| 1 | 4 | 0.7 | < 3 | 0.65 | 1 | 0.68 | 0.72 | 0.71 | 0.708 | 0.67 | [9 empty] | 0.69121
| 1 | 5 | 0.7 | < 3 | 0.65 | 0.8 | 0.64 | 0.76 | 0.71 | 0.687 | 0.71 | [9 empty] | 65
| 1 | 6 | 0.7 | < 3 | 0.64 | 0.8 | 0.68 | 0.72 | 0.68 | 0.684 | 0.68
| Mean |  | 0.7 |  | 0.65 | 0.73333 | 0.67333 | 0.735 | 0.70667 | 0.69633 | 0.68833
| Median |  | 0.7 |  | 0.65 | 0.7 | 0.67 | 0.725 | 0.71 | 0.6965 | 0.685
| Std Dev. |  | 1.2162e-16 |  | 0.0063246 | 0.1633 | 0.027325 | 0.019748 | 0.01633 | 0.014236 | 0.018348
| Rel.Std.Dev. |  | 1.737e-14% |  | 0.973% | 22.27% | 4.058% | 2.687% | 2.311% | 2.044% | 2.666%
| PDM3 |  | 1.272% |  | -5.962% | 6.094% | -2.586% | 6.336% | 2.236% | 0.7415% | -0.4159%
| Z-Score (Absolute) |  | 0 | 20.23 | 1.26 | 0.84 | 0.67 | 0.89 | 0.17 | 0.09 | 0.3
| Analytical results for Te in OREAS 926b (Indicative Value < 5 ppm) | [19 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 08 | 10 | 13 | 17 | 20 | 21 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 6 | < 2 | < 1 | < 5 | < 5 | 0.47631 | [12 empty] | 1
| 1 | 2 | < 6 | < 2 | < 1 | < 5 | < 5 | 0.49392 | [12 empty] | 12
| 1 | 3 | < 6 | < 2 | < 1 | < 5 | < 5 | 0.51735 | [12 empty] | 16
| 1 | 4 | < 6 | < 2 | < 1 | < 5 | < 5 | 0.56949 | [12 empty] | < 5
| 1 | 5 | < 6 | < 2 | < 1 | < 5 | < 5 | 0.54661 | [12 empty] | 18
| 1 | 6 | < 6 | < 2 | < 1 | < 5 | < 5 | 0.63056
| Mean | [6 empty] | 0.53904
| Median | [6 empty] | 0.53198
| Std Dev. | [6 empty] | 0.056218
| Rel.Std.Dev. | [6 empty] | 10.43%
PDM3
| Z-Score (Absolute) |  | 0.86 | 0.52 | 0.86 | 0.52 | 0.52 | 0.83
| Analytical results for Th in OREAS 926b (Certified Value 12.6 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | [7 empty] | 1
2
| 1 | 1 | 12.4 | 13.1 | < 50 | 13.8 | 13.5 | 13.3 | 12.57 | 12.8 | 11.07 | 12 | 11.536 | [7 empty] | 1
| 1 | 2 | 11.9 | 13.3 | < 50 | 13.7 | 14 | 12.5 | 12.41 | 13 | 11.254 | 12.7 | 11.6755 | [7 empty] | 5
| 1 | 3 | 12.4 | 13.3 | < 50 | 14 | 13.5 | 12.7 | 12.44 | 13 | 11.224 | 12.4 | 11.6137 | [7 empty] | 16
| 1 | 4 | 12.4 | 13 | < 50 | 13.7 | 14 | 12.7 | 12.4 | 13.5 | 11.272 | 12.4 | 11.3453 | [7 empty] | 12.6358
| 1 | 5 | 12.8 | 13.1 | < 50 | 13.7 | 13.5 | 12.1 | 12.59 | 13 | 11.295 | 12.8 | 11.5565 | [7 empty] | 66
| 1 | 6 | 12.8 | 13.3 | < 50 | 13.6 | 13.5 | 12.8 | 12.24 | 12.9 | 11.051 | 12.3 | 11.4033
| Mean |  | 12.45 | 13.1833 |  | 13.75 | 13.6667 | 12.6833 | 12.4417 | 13.0333 | 11.1943 | 12.4333 | 11.5217
| Median |  | 12.4 | 13.2 |  | 13.7 | 13.5 | 12.7 | 12.425 | 13 | 11.239 | 12.4 | 11.5463
| Std Dev. |  | 0.33317 | 0.13292 |  | 0.13784 | 0.2582 | 0.392 | 0.12797 | 0.24221 | 0.1064 | 0.28752 | 0.12545
| Rel.Std.Dev. |  | 2.676% | 1.008% |  | 1.002% | 1.889% | 3.091% | 1.029% | 1.858% | 0.9505% | 2.312% | 1.089%
| PDM3 |  | -1.47% | 4.333% |  | 8.818% | 8.159% | 0.3764% | -1.536% | 3.146% | -11.41% | -1.602% | -8.817%
| Z-Score (Absolute) |  | 0.31 | 0.67 | 16.61 | 1.44 | 1.33 | 0 | 0.33 | 0.47 | 2.01 | 0.34 | 1.57
| Analytical results for Ti in OREAS 926b (Certified Value 0.281 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | 3
3
| 1 | 1 | 0.27563 | 0.282 | 0.276 | 0.258 | 0.276 | 0.276 | 0.28 | 0.28 | 0.29 | 0.2873 | 0.29 | 0.22098 | 0.26 | 0.282 | 0.27 | 0.2828 | 0.28 | 0.26886 | 1
| 1 | 2 | 0.2908 | 0.276 | 0.276 | 0.258 | 0.276 | 0.27 | 0.29 | 0.28 | 0.29 | 0.285 | 0.28 | 0.22543 | 0.29 | 0.282 | 0.28 | 0.2712 | 0.29 | 0.28563 | 22
| 1 | 3 | 0.27826 | 0.288 | 0.276 | 0.258 | 0.276 | 0.258 | 0.29 | 0.28 | 0.29 | 0.3027 | 0.29 | 0.222 | 0.29 | 0.282 | 0.28 | 0.2754 | 0.28 | 0.27246 | 16
| 1 | 4 | 0.2831 | 0.282 | 0.276 | 0.312 | 0.276 | 0.27 | 0.28 | 0.28 | 0.29 | 0.2817 | 0.29 | 0.22463 | 0.29 | 0.282 | 0.28 | 0.28 | 0.28 | 0.27844 | 0.28116
| 1 | 5 | 0.28244 | 0.276 | 0.282 | 0.258 | 0.282 | 0.27 | 0.29 | 0.28 | 0.3 | 0.2894 | 0.29 | 0.22719 | 0.29 | 0.276 | 0.29 | 0.279 | 0.28 | 0.26287 | 67
| 1 | 6 | 0.28622 | 0.282 | 0.276 | 0.3 | 0.282 | 0.27 | 0.29 | 0.28 | 0.3 | 0.2985 | 0.29 | 0.22201 | 0.3 | 0.282 | 0.28 | 0.2808 | 0.29 | 0.27186
| Mean |  | 0.28274 | 0.281 | 0.277 | 0.274 | 0.278 | 0.269 | 0.28667 | 0.28 | 0.29333 | 0.29077 | 0.28833 | 0.22371 | 0.28667 | 0.281 | 0.28 | 0.2782 | 0.28333 | 0.27335
| Median |  | 0.28277 | 0.282 | 0.276 | 0.258 | 0.276 | 0.27 | 0.29 | 0.28 | 0.29 | 0.28835 | 0.29 | 0.22332 | 0.29 | 0.282 | 0.28 | 0.2795 | 0.28 | 0.27216
| Std Dev. |  | 0.0054384 | 0.0045166 | 0.0024495 | 0.025076 | 0.0030984 | 0.0058992 | 0.005164 | 0 | 0.005164 | 0.0081431 | 0.0040825 | 0.0024159 | 0.013663 | 0.0024495 | 0.0063246 | 0.0042105 | 0.005164 | 0.0078692
| Rel.Std.Dev. |  | 1.923% | 1.607% | 0.8843% | 9.152% | 1.115% | 2.193% | 1.801% | 0% | 1.76% | 2.801% | 1.416% | 1.08% | 4.766% | 0.8717% | 2.259% | 1.513% | 1.823% | 2.879%
| PDM3 |  | 0.5631% | -0.05669% | -1.479% | -2.546% | -1.124% | -4.325% | 1.959% | -0.4124% | 4.33% | 3.417% | 2.552% | -20.43% | 1.959% | -0.05669% | -0.4124% | -1.053% | 0.7732% | -2.776%
| Z-Score (Absolute) |  | 0.32 | 0.05 | 0.49 | 0.9 | 0.36 | 1.58 | 0.86 | 0.05 | 1.76 | 1.41 | 1.08 | 7.72 | 0.86 | 0.05 | 0.05 | 0.3 | 0.4 | 0.96
| Analytical results for Tl in OREAS 926b (Certified Value 0.97 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 10 | 13 | 15 | 17 | 18 | 20 | 21 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 1 | 1.1 | 1 | < 2 | 0.9 | 1 | 1 | 0.97564 | [10 empty] | 1
| 1 | 2 | 1 | 1.1 | 1.1 | < 2 | 1 | 1 | 1 | 0.93618 | [10 empty] | #N/A
| 1 | 3 | 0.9 | 1 | 1.2 | < 2 | 0.92 | 1 | 1 | 0.94899 | [10 empty] | 16
| 1 | 4 | 0.9 | 1 | 1 | < 2 | 1.05 | 1 | 0.9 | 0.96621 | [10 empty] | 0.97399
| 1 | 5 | 0.9 | 1.1 | 1.1 | < 2 | 0.92 | 1 | 1 | 0.97132 | [10 empty] | 68
| 1 | 6 | 1.1 | 1.1 | 1.1 | < 2 | 0.96 | 1 | 1 | 0.97145
| Mean |  | 0.96667 | 1.0667 | 1.0833 |  | 0.95833 | 1 | 0.98333 | 0.96163
| Median |  | 0.95 | 1.1 | 1.1 |  | 0.94 | 1 | 1 | 0.96876
| Std Dev. |  | 0.08165 | 0.05164 | 0.075277 |  | 0.057417 | 0 | 0.040825 | 0.01559
| Rel.Std.Dev. |  | 8.447% | 4.841% | 6.949% |  | 5.991% | 0% | 4.152% | 1.621%
| PDM3 |  | -0.7522% | 9.515% | 11.23% |  | -1.608% | 2.67% | 0.959% | -1.269%
| Z-Score (Absolute) |  | 0.61 | 1.84 | 2.25 | 0.2 | 0.82 | 0.2 | 0.2 | 0.74
| Analytical results for Tm in OREAS 926b (Certified Value 0.35 ppm) | [19 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.3 | < 2 | 0.35 | 0.2 | 0.35 | 0.33 | 0.34 | 0.362 | 0.38 | [9 empty] | 1
| 1 | 2 | 0.3 | < 2 | 0.34 | 0.4 | 0.32 | 0.37 | 0.35 | 0.358 | 0.38 | [9 empty] | 17
| 1 | 3 | 0.3 | < 2 | 0.33 | 0.2 | 0.32 | 0.36 | 0.38 | 0.373 | 0.38 | [9 empty] | 16
| 1 | 4 | 0.3 | < 2 | 0.33 | 0.6 | 0.33 | 0.33 | 0.37 | 0.366 | 0.36 | [9 empty] | 0.35483
| 1 | 5 | 0.3 | < 2 | 0.35 | 0.4 | 0.31 | 0.42 | 0.36 | 0.373 | 0.38 | [9 empty] | 69
| 1 | 6 | 0.3 | < 2 | 0.34 | 0.4 | 0.33 | 0.35 | 0.37 | 0.356 | 0.34
| Mean |  | 0.3 |  | 0.34 | 0.36667 | 0.32667 | 0.36 | 0.36167 | 0.36467 | 0.37
| Median |  | 0.3 |  | 0.34 | 0.4 | 0.325 | 0.355 | 0.365 | 0.364 | 0.38
| Std Dev. |  | 0 |  | 0.0089443 | 0.15055 | 0.013663 | 0.033466 | 0.01472 | 0.0073121 | 0.016733
| Rel.Std.Dev. |  | 0% |  | 2.631% | 41.06% | 4.182% | 9.296% | 4.07% | 2.005% | 4.522%
| PDM3 |  | -15.45% |  | -4.18% | 3.335% | -7.938% | 1.456% | 1.926% | 2.771% | 4.274%
| Z-Score (Absolute) |  | NA | 51.65 | 1.75 | NA | 2.83 | 0.13 | 0 | 0.24 | 0.67
Results from laboratories 8 and 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 926b (Certified Value 4.2 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | [7 empty] | 2
3
| 1 | 1 | 4.1 | 4.4 | < 50 | 4.2 | 4.5 | 4.3 | 3.91 | 4.41 | 3.512 | 4.1 | 4.0858 | [7 empty] | 1
| 1 | 2 | 3.9 | 4.5 | < 50 | 4.1 | 4.5 | 4.1 | 4.03 | 4.37 | 3.554 | 4.3 | 4.2916 | [7 empty] | #N/A
| 1 | 3 | 4 | 4.4 | < 50 | 4.1 | 4.5 | 4.1 | 4.04 | 4.39 | 3.541 | 4.2 | 4.3656 | [7 empty] | 16
| 1 | 4 | 4.1 | 4.1 | < 50 | 4 | 6 | 4.1 | 4.26 | 4.47 | 3.582 | 4.3 | 4.2962 | [7 empty] | 4.2017
| 1 | 5 | 4.1 | 4.5 | < 50 | 4.4 | 4.5 | 4 | 4.21 | 4.32 | 3.596 | 4.4 | 3.966 | [7 empty] | 70
| 1 | 6 | 4 | 4.7 | < 50 | 4.2 | 5 | 4.2 | 4.03 | 4.36 | 3.459 | 4.1 | 3.8752
| Mean |  | 4.0333 | 4.4333 |  | 4.1667 | 4.8333 | 4.1333 | 4.08 | 4.3867 | 3.5407 | 4.2333 | 4.1467
| Median |  | 4.05 | 4.45 |  | 4.15 | 4.5 | 4.1 | 4.035 | 4.38 | 3.5475 | 4.25 | 4.1887
| Std Dev. |  | 0.08165 | 0.19664 |  | 0.13663 | 0.60553 | 0.10328 | 0.13023 | 0.050859 | 0.049879 | 0.12111 | 0.20068
| Rel.Std.Dev. |  | 2.024% | 4.435% |  | 3.279% | 12.53% | 2.499% | 3.192% | 1.159% | 1.409% | 2.861% | 4.839%
| PDM3 |  | -4.006% | 5.514% |  | -0.8332% | 15.03% | -1.626% | -2.896% | 4.403% | -15.73% | 0.7535% | -1.308%
| Z-Score (Absolute) |  | 0.67 | 1.35 | 105.36 | 0 | 3.37 | 0.17 | 0.44 | 1.11 | 3.17 | 0.34 | 0.1
| Analytical results for V in OREAS 926b (Certified Value 177 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 171.7909 | 172 | 180 | 170 | 170 | 170 | 162.5395 | 150 | 189 | 178 | 165 | 188 | 179.6146 | [5 empty] | 1
| 1 | 2 | 181.2182 | 179 | 180 | 170 | 171 | 170 | 164.6295 | 150 | 188 | 180 | 168 | 188 | 181.932 | [5 empty] | 6
| 1 | 3 | 170.9545 | 171 | 174 | 170 | 178 | 170 | 166.7765 | 150 | 189 | 179 | 174 | 181 | 185.7953 | [5 empty] | 16
| 1 | 4 | 177.5545 | 175 | 178 | 170 | 171 | 170 | 163.312 | 150 | 188 | 184 | 172 | 184 | 185.6531 | [5 empty] | 176.5835
| 1 | 5 | 175.7545 | 168 | 180 | 170 | 175 | 170 | 166.6305 | 150 | 192 | 184 | 170 | 186 | 184.1997 | [5 empty] | 71
| 1 | 6 | 176.4727 | 172 | 175 | 180 | 186 | 170 | 168.882 | 150 | 190 | 184 | 171 | 188 | 185.3022
| Mean |  | 175.6242 | 172.8333 | 177.8333 | 171.6667 | 175.1667 | 170 | 165.4617 | 150 | 189.3333 | 181.5 | 170 | 185.8333 | 183.7495
| Median |  | 176.1136 | 172 | 179 | 170 | 173 | 170 | 165.63 | 150 | 189 | 182 | 170.5 | 187 | 184.7509
| Std Dev. |  | 3.8015 | 3.7639 | 2.7142 | 4.0825 | 6.1128 | 0 | 2.3936 | 0 | 1.5055 | 2.8107 | 3.1623 | 2.8577 | 2.4829
| Rel.Std.Dev. |  | 2.165% | 2.178% | 1.526% | 2.378% | 3.49% | 0% | 1.447% | 0% | 0.7952% | 1.549% | 1.86% | 1.538% | 1.351%
| PDM3 |  | -0.5432% | -2.124% | 0.7078% | -2.784% | -0.8024% | -3.728% | -6.298% | -15.05% | 7.22% | 2.784% | -3.728% | 5.238% | 4.058%
| Z-Score (Absolute) |  | 0.06 | 0.3 | 0.35 | 0.46 | 0 | 0.67 | 1.27 | 3.28 | 1.85 | 0.83 | 0.67 | 1.39 | 1.12
| Analytical results for W in OREAS 926b (Certified Value 6.5 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [7 empty] | 2
2
| 1 | 1 | 6.4 | 7 | < 100 | < 20 | 9 | 15 | 6.8 | 13 | < 100 | < 5 | 6.4512 | [7 empty] | 1
| 1 | 2 | 6.4 | 6 | < 100 | < 20 | 8 | 15 | 6.2 | 13 | < 100 | 5 | 6.4796 | [7 empty] | 7
| 1 | 3 | 6.1 | 6 | < 100 | < 20 | 9 | 15 | 6.3 | 13 | < 100 | 5 | 5.8597 | [7 empty] | 16
| 1 | 4 | 7.3 | 6 | < 100 | < 20 | 9 | 15 | 6.4 | 14 | < 100 | 5 | 5.6884 | [7 empty] | 6.5039
| 1 | 5 | 6.7 | 7 | < 100 | < 20 | 9 | 10 | 5.9 | 12 | < 100 | < 5 | 5.5411 | [7 empty] | 72
| 1 | 6 | 7.3 | 6 | < 100 | < 20 | 8 | 15 | 6.3 | 12 | < 100 | < 5 | 6.0204
| Mean |  | 6.7 | 6.3333 |  |  | 8.6667 | 14.1667 | 6.3167 | 12.8333 |  | 5 | 6.0067
| Median |  | 6.55 | 6 |  |  | 9 | 15 | 6.3 | 13 |  | 5 | 5.9401
| Std Dev. |  | 0.502 | 0.5164 |  |  | 0.5164 | 2.0412 | 0.29269 | 0.75277 |  | 0 | 0.39014
| Rel.Std.Dev. |  | 7.492% | 8.154% |  |  | 5.958% | 14.41% | 4.634% | 5.866% |  | 0% | 6.495%
| PDM3 |  | 3.015% | -2.623% |  |  | 33.25% | 117.8% | -2.879% | 97.32% |  | -23.12% | -7.644%
| Z-Score (Absolute) |  | 0.5 | 0.59 | 10.48 | 0.34 | 0 | 1.39 | 0.6 | 1.06 | 10.48 | 1.25 | 0.67
| Analytical results for Y in OREAS 926b (Certified Value 23.1 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | [7 empty] | 1
2
| 1 | 1 | 23.3 | 23.7 | 24 | 23.4 | 11.5019 | 23 | 24.7 | 21 | 22.5 | 22.326 | 23.3 | [7 empty] | 1
| 1 | 2 | 23.4 | 23.5 | 23 | 22.7 | 11.4703 | 23 | 23 | 22 | 22.48 | 22.724 | 23.4 | [7 empty] | #N/A
| 1 | 3 | 22 | 23.9 | 25 | 22.9 | 11.3732 | 23 | 23.3 | 22 | 23.32 | 23.583 | 22.9 | [7 empty] | 16
| 1 | 4 | 23.4 | 23.4 | 23 | 23.2 | 11.2742 | 23 | 22.5 | 22 | 22.76 | 23.403 | 22.9 | [7 empty] | 23.1075
| 1 | 5 | 23.5 | 24.2 | 24 | 22.4 | 11.4456 | 23 | 22.2 | 23 | 22.65 | 23.252 | 23.5 | [7 empty] | 73
| 1 | 6 | 23.4 | 24 | 24 | 22.3 | 11.8984 | 22 | 23.2 | 23 | 22.27 | 23.18 | 22.7
| Mean |  | 23.1667 | 23.7833 | 23.8333 | 22.8167 | 11.4939 | 22.8333 | 23.15 | 22.1667 | 22.6633 | 23.078 | 23.1167
| Median |  | 23.4 | 23.8 | 24 | 22.8 | 11.4579 | 23 | 23.1 | 22 | 22.575 | 23.216 | 23.1
| Std Dev. |  | 0.57504 | 0.30605 | 0.75277 | 0.43551 | 0.21422 | 0.40825 | 0.86891 | 0.75277 | 0.36214 | 0.46736 | 0.32506
| Rel.Std.Dev. |  | 2.482% | 1.287% | 3.158% | 1.909% | 1.864% | 1.788% | 3.753% | 3.396% | 1.598% | 2.025% | 1.406%
| PDM3 |  | 0.2562% | 2.925% | 3.141% | -1.258% | -50.26% | -1.186% | 0.1841% | -4.071% | -1.922% | -0.1275% | 0.03981%
| Z-Score (Absolute) |  | 0.23 | 1.82 | 1.95 | 0.67 | 29.89 | 0.63 | 0.19 | 2.35 | 1.07 | 0 | 0.1
| Analytical results for Yb in OREAS 926b (Certified Value 2.3 ppm) | [19 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 2.4 | 2 | 2.2 | 2.5 | 2.28 | 2.2 | 2.4 | 2.301 | 2.5 | [9 empty] | 1
| 1 | 2 | 2.3 | 2 | 2.3 | 2 | 2.17 | 2.5 | 2.3 | 2.326 | 2.4 | [9 empty] | 24
| 1 | 3 | 2.4 | 2 | 2.3 | 2 | 2.15 | 2.4 | 2.4 | 2.389 | 2.5 | [9 empty] | 16
| 1 | 4 | 2 | 2 | 2.2 | 7 | 2.2 | 2.5 | 2.4 | 2.355 | 2.4 | [9 empty] | 2.3043
| 1 | 5 | 2.2 | 2 | 2.2 | 3 | 2.09 | 2.6 | 2.5 | 2.413 | 2.5 | [9 empty] | 74
| 1 | 6 | 2.2 | 3 | 2.3 | 3 | 2.22 | 2.2 | 2.3 | 2.314 | 2.4
| Mean |  | 2.25 | 2.1667 | 2.25 | 3.25 | 2.185 | 2.4 | 2.3833 | 2.3497 | 2.45
| Median |  | 2.25 | 2 | 2.25 | 2.75 | 2.185 | 2.45 | 2.4 | 2.3405 | 2.45
| Std Dev. |  | 0.15166 | 0.40825 | 0.054772 | 1.8908 | 0.06473 | 0.16733 | 0.075277 | 0.044243 | 0.054772
| Rel.Std.Dev. |  | 6.74% | 18.84% | 2.434% | 58.18% | 2.962% | 6.972% | 3.158% | 1.883% | 2.236%
| PDM3 |  | -2.358% | -5.974% | -2.358% | 41.04% | -5.179% | 4.152% | 3.428% | 1.967% | 6.321%
| Z-Score (Absolute) |  | 0.79 | NA | 0.79 | 5.96 | 1.22 | 0.23 | 0.11 | 0.11 | 0.56
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 926b (Certified Value 110 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 0
0
| 1 | 1 | 115.8665 | 100 | 200 | 100 | 100 | 100 | 90 | 176 | 160 | 210 | 100 | 104.42 | 100 | 100 | 121 | 138 | 117 | 98.5403 | 1
| 1 | 2 | 113.4569 | 100 | 200 | 100 | 100 | 100 | 100 | 117 | 100 | 140 | 110 | 109.47 | 100 | 100 | 123 | 157 | 103 | 99.3279 | #N/A
| 1 | 3 | 121.8814 | 100 | 200 | 100 | 100 | 100 | 100 | 128 | 110 | 152 | 100 | 109.24 | 100 | < 100 | 125 | 113 | 100 | 93.5901 | 16
| 1 | 4 | 113.4754 | 100 | 200 | 100 | 100 | 100 | 100 | 105 | 110 | 127 | 100 | 108.49 | 100 | 100 | 121 | 119 | 102 | 104.3214 | 109.9748
| 1 | 5 | 113.8925 | 100 | 200 | 100 | 100 | 100 | 100 | 107 | 110 | 137 | 100 | 103.5 | 100 | 100 | 123 | 109 | 101 | 100.7906 | 75
| 1 | 6 | 115.1622 | 100 | 200 | 100 | 100 | 100 | 110 | 115 | 110 | 465 | 100 | 103.31 | 100 | < 100 | 124 | 102 | 104 | 107.4556
| Mean |  | 115.6225 | 100 | 200 | 100 | 100 | 100 | 100 | 124.6667 | 116.6667 | 205.1667 | 101.6667 | 106.405 | 100 | 100 | 122.8333 | 123 | 104.5 | 100.671
| Median |  | 114.5273 | 100 | 200 | 100 | 100 | 100 | 100 | 116 | 110 | 146 | 100 | 106.455 | 100 | 100 | 123 | 116 | 102.5 | 100.0592
| Std Dev. |  | 3.2168 | 0 | 0 | 0 | 0 | 0 | 6.3246 | 26.4474 | 21.6025 | 130.6651 | 4.0825 | 2.9576 | 0 | 0 | 1.6021 | 20.6591 | 6.2849 | 4.8136
| Rel.Std.Dev. |  | 2.782% | 0% | 0% | 0% | 0% | 0% | 6.325% | 21.21% | 18.52% | 63.69% | 4.016% | 2.78% | 0% | 0% | 1.304% | 16.8% | 6.014% | 4.781%
| PDM3 |  | 5.135% | -9.07% | 81.86% | -9.07% | -9.07% | -9.07% | -9.07% | 13.36% | 6.085% | 86.56% | -7.555% | -3.246% | -9.07% | -9.07% | 11.69% | 11.84% | -4.978% | -8.46%
| Z-Score (Absolute) |  | 0 | NA | NA | NA | NA | NA | NA | 0.66 | NA | 6.55 | 1.02 | 0.67 | NA | NA | 0.53 | 0.54 | 0.81 | 1.09
Results from laboratories 8, 11 and 15 were removed due to their 10 ppm reading resolution. Results from laboratories 3, 4, 5, 6, 7 and 16 were removed due to their 100 ppm reading resolution.
| Analytical results for Zr in OREAS 926b (Indicative Value 147 ppm) | [19 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 12 | 15 | 16 | 18 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*MS | [14 empty] | 0
0
| 1 | 1 | 125 | 150 | 167 | 153.1 | [14 empty] | 1
| 1 | 2 | 127 | 150 | 154.6 | 156.9 | [14 empty] | 9
| 1 | 3 | 137 | 150 | 153.1 | 167.3 | [14 empty] | 16
| 1 | 4 | 123 | 150 | 151.3 | 158.1 | [14 empty] | 146.8583
| 1 | 5 | 119 | 150 | 148.4 | 155.9 | [14 empty] | 19
| 1 | 6 | 116 | 150 | 154.6 | 157.3
| Mean |  | 124.5 | 150 | 154.8333 | 158.1
| Median |  | 124 | 150 | 153.85 | 157.1
| Std Dev. |  | 7.3144 | 0 | 6.4033 | 4.8283
| Rel.Std.Dev. |  | 5.875% | 0% | 4.136% | 3.054%
| PDM3 |  | -15.22% | 2.139% | 5.43% | 7.655%
| Z-Score (Absolute) |  | 4.65 | 0.4 | 0.4 | 0.95
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 926b (Certified Value 3.96 ppm) | [23 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 3.58 | 4.12 | 3.95 | 4.21 | 3.75 | 3.7 | 4.1 | 4.19 | 4.3 | 3.851 | 3.9 | 4.6 | 3.5275 | 3.8 | 4.1 | 3.93 | 4.09 | < 5 | 4.11 | 3.555 | 5.1 |  | 1
| 1 | 2 | 3.99 | 4.08 | 3.86 | 4.02 | 3.66 | 3.93 | 3.7 | 4.27 | 4.2 | 4.082 | 3.9 | 4.6 | 3.6822 | 3.8 | 4.2 | 3.94 | 3.95 | < 5 | 3.64 | 3.4991 | 4.46 |  | 11
| 1 | 3 | 3.82 | 4.19 | 3.94 | 4.08 | 3.7 | 3.92 | 4 | 4.07 | 4.2 | 4.035 | 4 | 4.4 | 3.5481 | 3.8 | 4.2 | 3.94 | 3.92 | < 5 | 3.86 | 3.4521 | 4.41 |  | 16
| 1 | 4 | 3.99 | 4.22 | 3.95 | 3.93 | 3.78 | 3.97 | 3.9 | 3.93 | 4 | 4.034 | 3.8 | 4.3 | 3.4347 | 3.8 | 4.2 | 3.93 | 3.79 | < 5 |  | 3.3056 | 4.29 |  | 3.9593
| 1 | 5 | 4.19 | 4.26 | 3.9 | 3.94 | 3.86 | 3.94 | 4.1 | 4.09 | 4 | 4.558 | 3.9 | 4.3 | 3.8266 | 3.8 | 4 | 3.7 | 3.9 | < 5 | 3.29 | 3.2943 | 4.35 |  | 77
| 1 | 6 | 3.46 | 4.28 | 4.01 | 3.84 | 3.88 | 4.08 | 4.1 | 3.98 | 4.2 | 4.184 | 3.9 | 4.6 | 3.5688 | 3.8 | 4.1 | 3.92 | 4.18 | < 5 | 3.15 | 3.4389 | 4.39
| Mean |  | 3.8383 | 4.1917 | 3.935 | 4.0033 | 3.7717 | 3.9233 | 3.9833 | 4.0883 | 4.15 | 4.124 | 3.9 | 4.4667 | 3.598 | 3.8 | 4.1333 | 3.8933 | 3.9717 |  | 3.61 | 3.4242 | 4.5
| Median |  | 3.905 | 4.205 | 3.945 | 3.98 | 3.765 | 3.935 | 4.05 | 4.08 | 4.2 | 4.0585 | 3.9 | 4.5 | 3.5585 | 3.8 | 4.15 | 3.93 | 3.935 |  | 3.64 | 3.4455 | 4.4
| Std Dev. |  | 0.27564 | 0.078592 | 0.050892 | 0.13033 | 0.086814 | 0.12404 | 0.16021 | 0.12687 | 0.12247 | 0.23844 | 0.063246 | 0.15055 | 0.13737 | 4.8648e-16 | 0.08165 | 0.095009 | 0.14049 |  | 0.39604 | 0.10456 | 0.29947
| Rel.Std.Dev. |  | 7.181% | 1.875% | 1.293% | 3.256% | 2.302% | 3.162% | 4.022% | 3.103% | 2.951% | 5.782% | 1.622% | 3.371% | 3.818% | 1.28e-14% | 1.975% | 2.44% | 3.537% |  | 10.97% | 3.054% | 6.655%
| PDM3 |  | -3.055% | 5.869% | -0.6139% | 1.112% | -4.739% | -0.9086% | 0.6068% | 3.259% | 4.816% | 4.16% | -1.498% | 12.81% | -9.126% | -4.024% | 4.395% | -1.666% | 0.3122% |  | -8.822% | -13.52% | 13.66%
| Z-Score (Absolute) |  | 0.4 | 1.06 | 0 | 0.28 | 0.67 | 0.05 | 0.2 | 0.63 | 0.89 | 0.78 | 0.14 | 2.19 | 1.39 | 0.56 | 0.82 | 0.17 | 0.15 | 5.92 | 1.34 | 2.11 | 2.33
| Analytical results for Al in OREAS 926b (Certified Value 5.69 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 5.8686 | 5.7068 | 5.53 | 5.63 | 5.71 | 5.72 | 5.56 | 5.35 | 5.77 | 5.85 | 5.64 | 5.6091 | 5.3 | 5.7393 | 5.97 | 5.85 | 5.46 | 4.41 | 5.6271 | 6 | 5.321 | 5.716 | 1
| 1 | 2 | 5.8341 | 5.6158 | 5.6 | 5.64 | 5.78 | 5.74 | 5.66 | 5.65 | 5.88 | 5.8 | 5.49 | 5.6511 | 5.33 | 5.7291 | 6.05 | 5.91 | 5.54 | 3.53 | 5.5348 | 5.86 | 5.296 | 5.716 | #N/A
| 1 | 3 | 5.894 | 5.6733 | 5.65 | 5.5 | 5.57 | 5.76 | 5.6 | 5.35 | 5.7 | 5.93 | 5.4 | 5.7587 | 5.44 | 5.7344 | 5.95 | 5.83 | 5.46 | 3.67 | 5.529 | 5.97 | 5.288 | 5.769 | 16
| 1 | 4 | 5.87 | 5.6231 | 5.6 | 5.7 | 5.64 | 5.75 | 5.68 | 5.98 | 5.7 | 5.84 | 6.01 | 5.5773 | 5.36 | 5.7864 | 5.91 | 5.76 | 5.65 | 3.86 | 5.6739 |  | 5.289 | 5.663 | 5.6932
| 1 | 5 | 5.8988 | 5.7067 | 5.74 | 5.58 | 5.72 | 5.72 | 5.66 | 5.44 | 5.73 | 5.69 | 5.81 | 5.6671 | 5.34 | 5.7364 | 6.03 | 5.75 | 5.35 | 4.82 | 5.4757 | 5.89 | 5.215 | 5.716 | 78
| 1 | 6 | 5.8265 | 5.6672 | 5.71 | 5.74 | 5.47 | 5.66 | 5.74 | 5.48 | 5.66 | 5.74 | 5.82 | 5.6292 | 5.54 | 5.7939 | 6 | 5.86 | 5.68 | 4.95 | 5.594 | 5.9 | 5.394 | 5.769
| Mean |  | 5.8653 | 5.6655 | 5.6383 | 5.6317 | 5.6483 | 5.725 | 5.65 | 5.5417 | 5.74 | 5.8083 | 5.695 | 5.6487 | 5.385 | 5.7533 | 5.985 | 5.8267 | 5.5233 | 4.2067 | 5.5724 | 5.924 | 5.3005 | 5.7248
| Median |  | 5.8693 | 5.6702 | 5.625 | 5.635 | 5.675 | 5.73 | 5.66 | 5.46 | 5.715 | 5.82 | 5.725 | 5.6402 | 5.35 | 5.7379 | 5.985 | 5.84 | 5.5 | 4.135 | 5.5644 | 5.9 | 5.2925 | 5.716
| Std Dev. |  | 0.029841 | 0.039334 | 0.077825 | 0.085421 | 0.11339 | 0.035637 | 0.062929 | 0.24145 | 0.077717 | 0.085186 | 0.2281 | 0.062435 | 0.089387 | 0.028874 | 0.052058 | 0.061536 | 0.12565 | 0.60599 | 0.072649 | 0.058566 | 0.057934 | 0.039897
| Rel.Std.Dev. |  | 0.5088% | 0.6943% | 1.38% | 1.517% | 2.007% | 0.6225% | 1.114% | 4.357% | 1.354% | 1.467% | 4.005% | 1.105% | 1.66% | 0.5019% | 0.8698% | 1.056% | 2.275% | 14.41% | 1.304% | 0.9886% | 1.093% | 0.6969%
| PDM3 |  | 3.023% | -0.4875% | -0.9642% | -1.081% | -0.7885% | 0.5581% | -0.7592% | -2.662% | 0.8216% | 2.022% | 0.03116% | -0.7812% | -5.414% | 1.054% | 5.125% | 2.344% | -2.984% | -26.11% | -2.122% | 4.053% | -6.898% | 0.5552%
| Z-Score (Absolute) |  | 1.55 | 0.06 | 0.14 | 0.19 | 0.07 | 0.5 | 0.06 | 0.87 | 0.61 | 1.12 | 0.28 | 0.07 | 2.03 | 0.71 | 2.44 | 1.26 | 1 | 10.82 | 0.64 | 1.99 | 2.66 | 0.5
| Analytical results for As in OREAS 926b (Certified Value 100 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 92 | 97.2 | 100.5 | 100 | 99.7 | 98.1 | 114 | 179 | 105 | 97.5 | 99.2 | 101 | 104 | 102 | 99 | 104 | 88.057 | 98.4 | 83.9359 |  |  |  | 1
| 1 | 2 | 90.78 | 96.7 | 99 | 104.5 | 98.9 | 97 | 104 | 188 | 106 | 94.2 | 99.2 | 102 | 103 | 101 | 99 | 99 | 88.936 | 95.8 | 86.8237 |  |  |  | 1
| 1 | 3 | 87.45 | 96.8 | 98.9 | 107 | 101 | 96 | 112 | 179 | 108 | 95.1 | 99.8 | 100 | 105 | 98 | 99 | 102 | 84.76 | 98.9 | 85.3707 |  |  |  | 16
| 1 | 4 | 93.46 | 98.5 | 100.5 | 105.5 | 99.8 | 98.4 | 104 | 178 | 102 | 98 | 97.6 | 101 | 103 | 100 | 94 | 97 | 89.755 |  | 83.9048 |  |  |  | 100.2843
| 1 | 5 | 92.53 | 99.2 | 100 | 105.5 | 99.1 | 99.7 | 103 | 183 | 101 | 98.6 | 101.6 | 101 | 103 | 101 | 94 | 105 | 88.349 | 97.9 | 80.6118 |  |  |  | 79
| 1 | 6 | 94.84 | 100.5 | 102.5 | 102.5 | 102 | 102 | 106 | 178 | 105 | 105.1 | 100.1 | 103 | 107 | 101 | 97 | 107 | 88.424 | 95.8 | 85.5053
| Mean |  | 91.8433 | 98.15 | 100.2333 | 104.1667 | 100.0833 | 98.5333 | 107.1667 | 180.8333 | 104.5 | 98.0833 | 99.5833 | 101.3333 | 104.1667 | 100.5 | 97 | 102.3333 | 88.0468 | 97.36 | 84.3587
| Median |  | 92.265 | 97.85 | 100.25 | 105 | 99.75 | 98.25 | 105 | 179 | 105 | 97.75 | 99.5 | 101 | 103.5 | 101 | 98 | 103 | 88.3865 | 97.9 | 84.6533
| Std Dev. |  | 2.5504 | 1.5215 | 1.314 | 2.5232 | 1.1923 | 2.1144 | 4.6655 | 3.9707 | 2.5884 | 3.8426 | 1.3121 | 1.0328 | 1.6021 | 1.3784 | 2.4495 | 3.7771 | 1.7173 | 1.4673 | 2.1372
| Rel.Std.Dev. |  | 2.777% | 1.55% | 1.311% | 2.422% | 1.191% | 2.146% | 4.353% | 2.196% | 2.477% | 3.918% | 1.318% | 1.019% | 1.538% | 1.372% | 2.525% | 3.691% | 1.95% | 1.507% | 2.534%
| PDM3 |  | -8.417% | -2.128% | -0.05081% | 3.871% | -0.2004% | -1.746% | 6.863% | 80.32% | 4.204% | -2.195% | -0.699% | 1.046% | 3.871% | 0.2151% | -3.275% | 2.043% | -12.2% | -2.916% | -15.88%
| Z-Score (Absolute) |  | 2.04 | 0.48 | 0.04 | 1.01 | 0 | 0.38 | 1.75 | 19.99 | 1.09 | 0.5 | 0.12 | 0.31 | 1.01 | 0.1 | 0.76 | 0.56 | 2.98 | 0.67 | 3.89
| Analytical results for B in OREAS 926b (Indicative Value 25.2 ppm) | [23 empty] | Indicative
B | Round | Replicate | Lab | [21 empty] | 1
| No | No | 17 | [21 empty] | ppm
|  |  | 4A*OES | [21 empty] | 1
1
| 1 | 1 | 27 | [21 empty] | 1
| 1 | 2 | 23 | [21 empty] | 16
| 1 | 3 | 26 | [21 empty] | 16
| 1 | 4 | 26 | [21 empty] | 25.1667
| 1 | 5 | 23 | [21 empty] | 22
| 1 | 6 | 26
| Mean |  | 25.1667
| Median |  | 26
| Std Dev. |  | 1.7224
| Rel.Std.Dev. |  | 6.844%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 926b (Certified Value 0.168 wt.%) | [23 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.16801 | 0.142 | 0.168 | 0.093 | 0.068 | 0.17 | 0.0104 | 0.156 | 0.1721 | 0.0131 | 0.1777 | 0.165 |  | 0.173 | 0.1368 | 0.1597 | 0.1678 | 0.17102 | 0.0505 | 0.158 | 0.181 |  | 1
| 1 | 2 | 0.16545 | 0.118 | 0.167 | 0.116 | 0.072 | 0.161 | 0.0192 | 0.164 | 0.1691 | 0.0105 | 0.1776 | 0.162 |  | 0.174 | 0.1357 | 0.1635 | 0.1611 | 0.16818 | 0.0663 | 0.158 | 0.169 |  | 20
| 1 | 3 | 0.16752 | 0.113 | 0.159 | 0.056 | 0.07 | 0.168 | 0.0164 | 0.16 | 0.1747 | 0.0136 | 0.1806 | 0.164 |  | 0.172 | 0.1343 | 0.1626 | 0.169 | 0.16763 | 0.0814 | 0.157 | 0.182 |  | 16
| 1 | 4 | 0.16544 | 0.096 | 0.162 | 0.083 | 0.071 | 0.168 | 0.0234 | 0.156 | 0.1718 | 0.0158 | 0.1745 | 0.167 |  | 0.171 | 0.1457 | 0.1603 | 0.1691 | 0.17095 |  | 0.155 | 0.169 |  | 0.16785
| 1 | 5 | 0.16782 | 0.101 | 0.151 | 0.07 | 0.109 | 0.171 | 0.0241 | 0.157 | 0.1694 | 0.0178 | 0.1789 | 0.165 |  | 0.172 | 0.1422 | 0.1536 | 0.1826 | 0.16584 | 0.207 | 0.149 | 0.171 |  | 80
| 1 | 6 | 0.16608 | 0.1 | 0.156 | 0.078 | 0.097 | 0.174 | 0.0176 | 0.157 | 0.1716 | 0.0151 | 0.1773 | 0.165 |  | 0.178 | 0.1351 | 0.1776 | 0.1863 | 0.16992 | 0.202 | 0.153 | 0.172
| Mean |  | 0.16672 | 0.11167 | 0.1605 | 0.082667 | 0.081167 | 0.16867 | 0.018517 | 0.15833 | 0.17145 | 0.014317 | 0.17777 | 0.16467 |  | 0.17333 | 0.1383 | 0.16288 | 0.17265 | 0.16892 | 0.12144 | 0.155 | 0.174
| Median |  | 0.1668 | 0.107 | 0.1605 | 0.0805 | 0.0715 | 0.169 | 0.0184 | 0.157 | 0.1717 | 0.01435 | 0.17765 | 0.165 |  | 0.1725 | 0.13625 | 0.16145 | 0.16905 | 0.16905 | 0.0814 | 0.156 | 0.1715
| Std Dev. |  | 0.0011994 | 0.017072 | 0.0065345 | 0.020549 | 0.017383 | 0.0043665 | 0.0050337 | 0.0031411 | 0.0020423 | 0.0025103 | 0.0020116 | 0.001633 |  | 0.0025033 | 0.0045874 | 0.0080004 | 0.0096765 | 0.0020606 | 0.076627 | 0.0035214 | 0.005933
| Rel.Std.Dev. |  | 0.7194% | 15.29% | 4.071% | 24.86% | 21.42% | 2.589% | 27.18% | 1.984% | 1.191% | 17.53% | 1.132% | 0.9917% |  | 1.444% | 3.317% | 4.912% | 5.605% | 1.22% | 63.1% | 2.272% | 3.41%
| PDM3 |  | -0.6714% | -33.47% | -4.377% | -50.75% | -51.64% | 0.489% | -88.97% | -5.667% | 2.147% | -91.47% | 5.911% | -1.894% |  | 3.269% | -17.6% | -2.957% | 2.862% | 0.6397% | -27.65% | -7.653% | 3.667%
| Z-Score (Absolute) |  | 0.3 | 2.99 | 0.07 | 4.72 | 4.81 | 0.42 | 8.54 | 0.2 | 0.58 | 8.79 | 0.96 | 0.18 | NA | 0.69 | 1.4 | 0.07 | 0.65 | 0.43 | 2.4 | 0.4 | 0.73
| Analytical results for Be in OREAS 926b (Certified Value 1.99 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.97 | 2.01 | 1.96 | 1.99 | 1.95 | 2.2 | 2.16 | 2.1 | 2 | 1.9 | 2 | 1.5175 | 1.5 | < 5 | 2.08 | 2.1 | < 5 | 2.46 | 2.2296 | 2.5 |  |  | 1
| 1 | 2 | 1.94 | 2 | 2.04 | 1.99 | 1.99 | 2 | 1.51 | 1.8 | 3 | 1.9 | 2 | 1.4554 | 1.5 | < 5 | 2.16 | 2.1 | < 5 | 2.21 | 2.3119 | 2.2 |  |  | #N/A
| 1 | 3 | 1.96 | 1.95 | 1.97 | 2.02 | 1.95 | 2.2 | 1.91 | 2.1 | 3 | 1.94 | 1.9 | 1.4672 | 1.5 | < 5 | 2.21 | 2.1 | < 5 | 2.18 | 2.0943 | 2.3 |  |  | 16
| 1 | 4 | 1.95 | 2.01 | 1.99 | 2 | 1.95 | 2 | 1.59 | 1.9 | 3 | 1.87 | 2 | 1.4646 | 1.5 | < 5 | 2.08 | 2 | < 5 |  | 2.0082 | 2.3 |  |  | 1.9926
| 1 | 5 | 2.05 | 1.97 | 2.01 | 1.99 | 1.97 | 1.8 | 1.7 | 1.8 | 3 | 1.93 | 1.9 | 1.4569 | 2 | < 5 | 2.07 | 2 | < 5 | 2.44 | 1.9894 | 2.2 |  |  | 81
| 1 | 6 | 1.97 | 2.04 | 1.94 | 1.99 | 2.04 | 1.8 | 1.71 | 1.9 | 2 | 1.9 | 2 | 1.3637 | 2 | < 5 | 2.05 | 2.2 | < 5 | 2.1 | 2.0161 | 2.2
| Mean |  | 1.9733 | 1.9967 | 1.985 | 1.9967 | 1.975 | 2 | 1.7633 | 1.9333 | 2.6667 | 1.9067 | 1.9667 | 1.4542 | 1.6667 |  | 2.1083 | 2.0833 |  | 2.278 | 2.1082 | 2.2833
| Median |  | 1.965 | 2.005 | 1.98 | 1.99 | 1.96 | 2 | 1.705 | 1.9 | 3 | 1.9 | 2 | 1.4607 | 1.5 |  | 2.08 | 2.1 |  | 2.21 | 2.0552 | 2.25
| Std Dev. |  | 0.039328 | 0.032042 | 0.036194 | 0.012111 | 0.035637 | 0.17889 | 0.23662 | 0.13663 | 0.5164 | 0.025033 | 0.05164 | 0.049974 | 0.2582 |  | 0.062423 | 0.075277 |  | 0.16223 | 0.13342 | 0.1169
| Rel.Std.Dev. |  | 1.993% | 1.605% | 1.823% | 0.6065% | 1.804% | 8.944% | 13.42% | 7.067% | 19.36% | 1.313% | 2.626% | 3.437% | 15.49% |  | 2.961% | 3.613% |  | 7.122% | 6.329% | 5.12%
| PDM3 |  | -0.9648% | 0.2062% | -0.3793% | 0.2062% | -0.8811% | 0.3735% | -11.5% | -2.972% | 33.83% | -4.311% | -1.299% | -27.02% | -16.36% |  | 5.81% | 4.556% |  | 14.33% | 5.806% | 14.59%
| Z-Score (Absolute) |  | 0.17 | 0 | 0.09 | 0 | 0.16 | 0.02 | 1.75 | 0.47 | NA | 0.67 | 0.22 | 4.06 | 2.47 | 3.77 | 0.84 | 0.65 | 3.77 | 2.11 | 0.84 | 2.15
| Analytical results for Bi in OREAS 926b (Certified Value 48 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 38.0531 | 46.9875 | 47.7 | 54.7 | 51.3 | 47.1 | 49.3 | 49 | 46.9 | 50.6 | 49.02 | 45.41 | 50.4 | 36.225 | 51.9 | 51.4 | 48 | 45.59 | 33.391 | 49.7 | 64.077 | 44.5 | 1
| 1 | 2 | 38.6107 | 46.025 | 49.4 | 53.6 | 47.4 | 46.5 | 50 | 44.4 | 49 | 50.2 | 51.04 | 44.74 | 51.9 | 36.321 | 51.1 | 49.1 | 47 | 42.95 | 35.443 | 49.3 | 63.0996 | 44 | 2
| 1 | 3 | 40.7066 | 45.2625 | 49 | 53.8 | 45.5 | 46.4 | 49.9 | 45.7 | 46.3 | 50.4 | 49.16 | 49.11 | 51.9 | 36.5647 | 52.2 | 49.8 | 46 | 44.38 | 30.692 | 51.8 | 63.5265 | 44.3 | 16
| 1 | 4 | 38.0074 | 45.65 | 48.3 | 55.5 | 45.4 | 47.4 | 49.6 | 45.2 | 45.7 | 49.8 | 50.52 | 47.29 | 51.4 | 36.453 | 50.7 | 49.6 | 47 | 43.4 | 29.167 |  | 64.5002 | 43.5 | 48.0083
| 1 | 5 | 39.509 | 48.9 | 50 | 54 | 45.3 | 47.6 | 49.7 | 44.4 | 46.5 | 50.1 | 53.96 | 48.61 | 51.4 | 36.5157 | 51.4 | 50.5 | 47 | 47.16 | 27.078 | 48 | 39.9425 | 43.7 | 82
| 1 | 6 | 41.1497 | 45.1875 | 48.1 | 55.5 | 44.6 | 47.9 | 49.8 | 44.7 | 46.5 | 50.5 | 50.35 | 46.75 | 52 | 36.6979 | 51.8 | 50.9 | 49 | 49.2 | 34.277 | 48.5 | 62.5424 | 43.9
| Mean |  | 39.3394 | 46.3354 | 48.75 | 54.5167 | 46.5833 | 47.15 | 49.7167 | 45.5667 | 46.8167 | 50.2667 | 50.675 | 46.985 | 51.5 | 36.4629 | 51.5167 | 50.2167 | 47.3333 | 45.4467 | 31.6747 | 49.46 | 59.6147 | 43.9833
| Median |  | 39.0598 | 45.8375 | 48.65 | 54.35 | 45.45 | 47.25 | 49.75 | 44.95 | 46.5 | 50.3 | 50.435 | 47.02 | 51.65 | 36.4844 | 51.6 | 50.15 | 47 | 44.985 | 32.0415 | 49.3 | 63.313 | 43.95
| Std Dev. |  | 1.3516 | 1.4168 | 0.86891 | 0.84715 | 2.4927 | 0.60249 | 0.24833 | 1.7558 | 1.1392 | 0.29439 | 1.7938 | 1.7219 | 0.6 | 0.1704 | 0.55648 | 0.86583 | 1.0328 | 2.3918 | 3.2342 | 1.4673 | 9.6622 | 0.37103
| Rel.Std.Dev. |  | 3.436% | 3.058% | 1.782% | 1.554% | 5.351% | 1.278% | 0.4995% | 3.853% | 2.433% | 0.5857% | 3.54% | 3.665% | 1.165% | 0.4673% | 1.08% | 1.724% | 2.182% | 5.263% | 10.21% | 2.967% | 16.21% | 0.8436%
| PDM3 |  | -18.06% | -3.485% | 1.545% | 13.56% | -2.968% | -1.788% | 3.558% | -5.086% | -2.482% | 4.704% | 5.555% | -2.132% | 7.273% | -24.05% | 7.308% | 4.6% | -1.406% | -5.336% | -34.02% | 3.024% | 24.18% | -8.384%
| Z-Score (Absolute) |  | 1.96 | 0.22 | 0.37 | 1.8 | 0.16 | 0.02 | 0.61 | 0.41 | 0.11 | 0.75 | 0.85 | 0.06 | 1.05 | 2.67 | 1.06 | 0.74 | 0.02 | 0.44 | 3.85 | 0.55 | 3.06 | 0.81
| Analytical results for Ca in OREAS 926b (Certified Value 0.038 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.044846 | 0.032383 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | 0.04 | 0.0566 | 0.032 | 0.036235 | 0.04 | 0.03 | 0.03 | 0.03 | 0.049535 | 0.05 | 0.048 | 0.036 | 1
| 1 | 2 | 0.042946 | 0.03245 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.0547 | 0.032 | 0.035679 | 0.04 | 0.03 | 0.03 | 0.03 | 0.049879 | 0.05 | 0.046 | 0.036 | 12
| 1 | 3 | 0.043294 | 0.03161 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.0551 | 0.033 | 0.035812 | 0.04 | 0.03 | 0.04 | 0.03 | 0.046499 | 0.05 | 0.043 | 0.036 | 16
| 1 | 4 | 0.043467 | 0.03108 | 0.03 | 0.03 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.04 | 0.0547 | 0.032 | 0.036193 | 0.04 | 0.03 | 0.04 | 0.03 | 0.068534 |  | 0.044 | 0.036 | 0.037763
| 1 | 5 | 0.04504 | 0.03135 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.05 | 0.04 | 0.0863 | 0.031 | 0.035846 | 0.04 | 0.03 | 0.03 | 0.03 | 0.042081 | 0.04 | 0.043 | 0.036 | 83
| 1 | 6 | 0.045895 | 0.032093 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.05 | 0.04 | 0.054 | 0.034 | 0.036479 | 0.04 | 0.03 | 0.04 | 0.03 | 0.038681 | 0.04 | 0.045 | 0.036
| Mean |  | 0.044248 | 0.031828 | 0.038333 | 0.033333 | 0.036667 | 0.033333 | 0.04 | 0.033333 | 0.041667 | 0.045 | 0.04 | 0.060233 | 0.032333 | 0.036041 | 0.04 | 0.03 | 0.035 | 0.03 | 0.049202 | 0.046 | 0.044833 | 0.036
| Median |  | 0.044156 | 0.031852 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.045 | 0.04 | 0.0549 | 0.032 | 0.036019 | 0.04 | 0.03 | 0.035 | 0.03 | 0.048017 | 0.05 | 0.0445 | 0.036
| Std Dev. |  | 0.0011757 | 0.0005658 | 0.0040825 | 0.005164 | 0.005164 | 0.005164 | 0 | 0.005164 | 0.0040825 | 0.0054772 | 0 | 0.012799 | 0.0010328 | 0.00030793 | 0 | 0 | 0.0054772 | 0 | 0.010421 | 0.0054772 | 0.0019408 | 0
| Rel.Std.Dev. |  | 2.657% | 1.778% | 10.65% | 15.49% | 14.08% | 15.49% | 0% | 15.49% | 9.798% | 12.17% | 0% | 21.25% | 3.194% | 0.8544% | 0% | 0% | 15.65% | 0% | 21.18% | 11.91% | 4.329% | 0%
| PDM3 |  | 17.17% | -15.72% | 1.511% | -11.73% | -2.903% | -11.73% | 5.924% | -11.73% | 10.34% | 19.16% | 5.924% | 59.5% | -14.38% | -4.56% | 5.924% | -20.56% | -7.316% | -20.56% | 30.29% | 21.81% | 18.72% | -4.668%
| Z-Score (Absolute) |  | 1.09 | 0.92 | 0.13 | 0.67 | 0.13 | 0.67 | 0.4 | 0.67 | 0.67 | 1.21 | 0.4 | 3.68 | 0.84 | 0.24 | 0.4 | 1.21 | 0.4 | 1.21 | 1.89 | 1.38 | 1.19 | 0.29
| Analytical results for Cd in OREAS 926b (Certified Value 0.16 ppm) | [23 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | < 5 | < 5 | 0.251 | 0.16 | 0.14 | 0.166 | 0.17 | < 0.3 | 0.49 | 0.4 | 0.17 | 0.29 | 0.2 | < 0.5 | < 0.5 | 0.29 | 0.21 | < 5 | 0.28 | 0.22505 | 0.8 |  | 1
| 1 | 2 | < 5 | < 5 | 0.147 | 0.12 | 0.15 | 0.139 | 0.15 | 0.9 | 0.25 | 0.2 | 0.06 | 0.17 | 0.21 | < 0.5 | < 0.5 | 0.28 | 0.14 | < 5 | 0.21 | 0.23587 | 0.4 |  | #N/A
| 1 | 3 | < 5 | < 5 | 0.134 | 0.13 | 0.15 | 0.126 | 0.16 | < 0.3 | 0.4 | 0.3 | 0.11 | 0.18 | 0.18 | < 0.5 | < 0.5 | 0.28 | 0.18 | < 5 | 0.18 | 0.20219 | 0.2 |  | 16
| 1 | 4 | < 5 | < 5 | 0.154 | 0.16 | 0.14 | 0.138 | 0.15 | 0.9 | 0.18 | < 0.2 | 0.16 | 0.17 | 0.19 | < 0.5 | < 0.5 | 0.3 | 0.16 | < 5 |  | 0.19538 | 0.2 |  | 0.16255
| 1 | 5 | < 5 | < 5 | 0.153 | 0.18 | 0.14 | 0.135 | 0.17 | < 0.3 | 0.23 | < 0.2 | 0.1 | 0.18 | 0.19 | < 0.5 | < 0.5 | 0.3 | 0.14 | < 5 | 0.16 | 0.20051 | 0.3 |  | 84
| 1 | 6 | < 5 | < 5 | 0.069 | 0.15 | 0.15 | 0.135 | 0.12 | < 0.3 | 0.29 | 0.3 | 0.08 | 0.19 | 0.18 | < 0.5 | < 0.5 | 0.28 | 0.14 | < 5 | 0.16 | 0.20745 | 0.2
| Mean |  |  |  | 0.15133 | 0.15 | 0.145 | 0.13983 | 0.15333 | 0.9 | 0.30667 | 0.3 | 0.11333 | 0.19667 | 0.19167 |  |  | 0.28833 | 0.16167 |  | 0.198 | 0.21108 | 0.35
| Median |  |  |  | 0.15 | 0.155 | 0.145 | 0.1365 | 0.155 | 0.9 | 0.27 | 0.3 | 0.105 | 0.18 | 0.19 |  |  | 0.285 | 0.15 |  | 0.18 | 0.20482 | 0.25
| Std Dev. |  |  |  | 0.05838 | 0.021909 | 0.0054772 | 0.013615 | 0.018619 | 0 | 0.11639 | 0.08165 | 0.043665 | 0.046332 | 0.01169 |  |  | 0.0098319 | 0.028577 |  | 0.0502 | 0.015875 | 0.23452
| Rel.Std.Dev. |  |  |  | 38.58% | 14.61% | 3.777% | 9.737% | 12.14% | 0% | 37.95% | 27.22% | 38.53% | 23.56% | 6.099% |  |  | 3.41% | 17.68% |  | 25.35% | 7.521% | 67.01%
| PDM3 |  |  |  | -6.898% | -7.719% | -10.79% | -13.97% | -5.668% | 453.7% | 88.66% | 84.56% | -30.28% | 20.99% | 17.92% |  |  | 77.39% | -0.5411% |  | 21.81% | 29.86% | 115.3%
| Z-Score (Absolute) |  | 29.58 | 29.58 | 0.59 | 0.61 | 0.67 | 0.74 | 0.57 | NA | 1.4 | NA | 1.08 | 0.01 | 0.07 | 0.68 | 0.68 | 1.17 | 0.46 | 29.58 | 0.01 | 0.18 | NA
Results from laboratories 8, 10 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 926b (Certified Value 67 ppm) | [23 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 0
1
| 1 | 1 | 71.8 | 65.1 | 70.8 | 69.8 | 69.5 | 73.5 | 54.5 | 69.5 | 65.47 | 65.3 | 59.9376 | 71.5 | 74 | 67 | 66.19 | 71.8 | 62.5993 | 62 |  |  |  |  | 1
| 1 | 2 | 72.3 | 66.2 | 69.3 | 68.2 | 69 | 69.8 | 59.6 | 67.8 | 65.32 | 65.1 | 59.5434 | 69.9 | 72 | 68 | 61.62 | 68.5 | 63.5004 | 61 |  |  |  |  | #N/A
| 1 | 3 | 73.2 | 67.7 | 72.3 | 69.5 | 67 | 69.7 | 57.8 | 63.3 | 65.14 | 66.8 | 59.4704 | 70.8 | 72 | 70 | 65.73 | 70.9 | 63.9475 | 63 |  |  |  |  | 16
| 1 | 4 | 73.9 | 68.3 | 70.7 | 70.9 | 70.6 | 69 | 58.1 | 64.2 | 64.82 | 64.2 | 59.0973 | 70.2 | 72 | 69 | 62.81 |  | 63.3306 | 62 |  |  |  |  | 67.3413
| 1 | 5 | 74.5 | 65.3 | 70.8 | 70.4 | 70.7 | 69.4 | 59.3 | 62.5 | 70.26 | 66.3 | 59.6675 | 70 | 72 | 68 | 67.95 | 78.5 | 59.7468 | 62 |  |  |  |  | 85
| 1 | 6 | 75 | 69 | 69.1 | 71.6 | 71.7 | 70.7 | 58.7 | 65.2 | 65.35 | 66.2 | 59.4823 | 71.8 | 74 | 67 | 73.3 | 74.3 | 63.1261 | 61
| Mean |  | 73.45 | 66.9333 | 70.5 | 70.0667 | 69.75 | 70.35 | 58 | 65.4167 | 66.06 | 65.65 | 59.5331 | 70.7 | 72.6667 | 68.1667 | 66.2667 | 72.8 | 62.7085 | 61.8333
| Median |  | 73.55 | 66.95 | 70.75 | 70.1 | 70.05 | 69.75 | 58.4 | 64.7 | 65.335 | 65.75 | 59.5129 | 70.5 | 72 | 68 | 65.96 | 71.8 | 63.2283 | 62
| Std Dev. |  | 1.2502 | 1.6305 | 1.1713 | 1.186 | 1.6502 | 1.6429 | 1.8461 | 2.7154 | 2.07 | 0.95656 | 0.27502 | 0.80498 | 1.0328 | 1.169 | 4.1506 | 3.8026 | 1.5171 | 0.75277
| Rel.Std.Dev. |  | 1.702% | 2.436% | 1.661% | 1.693% | 2.366% | 2.335% | 3.183% | 4.151% | 3.133% | 1.457% | 0.462% | 1.139% | 1.421% | 1.715% | 6.263% | 5.223% | 2.419% | 1.217%
| PDM3 |  | 9.071% | -0.6059% | 4.691% | 4.047% | 3.577% | 4.468% | -13.87% | -2.858% | -1.903% | -2.512% | -11.6% | 4.988% | 7.908% | 1.226% | -1.596% | 8.106% | -6.88% | -8.179%
| Z-Score (Absolute) |  | 1.38 | 0.14 | 0.69 | 0.59 | 0.52 | 0.66 | 2.24 | 0.5 | 0.35 | 0.45 | 1.88 | 0.74 | 1.2 | 0.14 | 0.3 | 1.23 | 1.14 | 1.34
| Analytical results for Co in OREAS 926b (Certified Value 7.01 ppm) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 11.2198 | 11.91 | 6.61 | 7.2 | 7.4 | 6.84 | 6.8 | 7 | 7.8 | 8 | 7.4 | 7.9 | 7 | 6.1145 | 6 | 7 | 7.4 | 7.3 | 6.93 | 6.37 | 6.8987 | 7.44 | 1
| 1 | 2 | 10.7672 | 10.88 | 6.63 | 6.6 | 7.4 | 6.67 | 6.5 | 7 | 7.2 | 7 | 7.5 | 7.9 | 7 | 6.0839 | 6 | 7 | 7.3 | 6.9 | 6.56 | 6.07 | 7.4123 | 7.17 | 4
| 1 | 3 | 11.4397 | 12.41 | 6.8 | 6.5 | 6.8 | 6.7 | 6.3 | 6 | 6.9 | 7 | 7.5 | 8.2 | 7 | 6.2873 | 6 | 7 | 7.5 | 6.9 | 7.09 | 6.28 | 6.7615 | 7.54 | 16
| 1 | 4 | 10.8534 | 10.88 | 6.93 | 6.6 | 6.5 | 6.82 | 6.5 | 7 | 7 | 7 | 7.8 | 8.1 | 7 | 6.2075 | 6 | 7 | 7.2 | 6.7 | 6.38 |  | 6.9392 | 7.39 | 7.0075
| 1 | 5 | 11.3276 | 10.9 | 6.91 | 6.6 | 6.8 | 6.83 | 6.6 | 7 | 7.9 | 7 | 8.2 | 15.4 | 7 | 6.195 | 6 | 7 | 7.2 | 7 | 6.73 | 7.54 | 6.4616 | 7.28 | 86
| 1 | 6 | 11.3793 | 10.74 | 7.03 | 6.7 | 6.4 | 6.95 | 6.7 | 7 | 7.3 | 7 | 7.7 | 7.8 | 7 | 6.277 | 6 | 7 | 7.3 | 7.4 | 6.95 | 7.57 | 6.8774 | 7.3
| Mean |  | 11.1645 | 11.2867 | 6.8183 | 6.7 | 6.8833 | 6.8017 | 6.5667 | 6.8333 | 7.35 | 7.1667 | 7.6833 | 9.2167 | 7 | 6.1942 | 6 | 7 | 7.3167 | 7.0333 | 6.7733 | 6.766 | 6.8918 | 7.3533
| Median |  | 11.2737 | 10.89 | 6.855 | 6.6 | 6.8 | 6.825 | 6.55 | 7 | 7.25 | 7 | 7.6 | 8 | 7 | 6.2012 | 6 | 7 | 7.3 | 6.95 | 6.83 | 6.37 | 6.888 | 7.345
| Std Dev. |  | 0.28498 | 0.69707 | 0.17023 | 0.25298 | 0.43089 | 0.10226 | 0.17512 | 0.40825 | 0.41352 | 0.40825 | 0.29269 | 3.0328 | 0 | 0.082749 | 0 | 0 | 0.1169 | 0.26583 | 0.26763 | 0.72851 | 0.30838 | 0.13079
| Rel.Std.Dev. |  | 2.553% | 6.176% | 2.497% | 3.776% | 6.26% | 1.503% | 2.667% | 5.974% | 5.626% | 5.696% | 3.809% | 32.91% | 0% | 1.336% | 0% | 0% | 1.598% | 3.78% | 3.951% | 10.77% | 4.475% | 1.779%
| PDM3 |  | 59.32% | 61.07% | -2.699% | -4.388% | -1.771% | -2.937% | -6.29% | -2.485% | 4.888% | 2.272% | 9.645% | 31.53% | -0.1065% | -11.61% | -14.38% | -0.1065% | 4.412% | 0.3692% | -3.341% | -3.446% | -1.651% | 4.936%
| Z-Score (Absolute) |  | 8.86 | 9.12 | 0.15 | 0.4 | 0.02 | 0.19 | 0.67 | NA | 0.95 | NA | 1.64 | 4.82 | NA | 1.45 | NA | NA | 0.88 | 0.29 | 0.25 | 0.26 | 0 | 0.96
Results from laboratories 8, 10, 13, 15 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 926b (Certified Value 74 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 0
1
| 1 | 1 | 64.726 | 79.1648 | 73 | 73 | 78 | 72.6 | 72 | 75 | 73 | 80 | 92 | 77.9 | 74 | 71.242 | 60 | 65 | 83 | 51 | 91.688 | 63.7 | 45.8229 | 83 | 1
| 1 | 2 | 63.0008 | 78.3952 | 73.7 | 75 | 78 | 72.9 | 71 | 73 | 75 | 80 | 89 | 78.3 | 76 | 71.415 | 70 | 65 | 78 | 51 | 90.148 | 61.1 | 56.2752 | 81 | 5
| 1 | 3 | 64.4156 | 78.5928 | 74.3 | 74 | 74 | 73.9 | 70 | 77 | 73 | 82 | 93 | 78.9 | 75 | 71.4635 | 70 | 64 | 77 | 52 | 90.947 | 63.3 | 47.4212 | 83 | 16
| 1 | 4 | 64.0166 | 77.8232 | 76.2 | 75 | 75 | 74.6 | 69 | 79 | 72 | 80 | 96 | 76.3 | 71 | 71.327 | 60 | 65 | 80 | 54 | 91.47 |  | 45.3098 | 84 | 73.879
| 1 | 5 | 62.927 | 79.82 | 76.4 | 73 | 80 | 75.3 | 72 | 69 | 74 | 76 | 95 | 78.7 | 70 | 71.891 | 60 | 67 | 77 | 53 | 88.417 | 57.3 | 51.3246 | 82 | 87
| 1 | 6 | 62.9859 | 79.2064 | 76.1 | 77 | 75 | 73.7 | 72 | 83 | 73 | 78 | 99 | 77.8 | 76 | 71.8425 | 60 | 67 | 77 | 47 | 91.547 | 73.6 | 46.9931 | 83
| Mean |  | 63.6787 | 78.8337 | 74.95 | 74.5 | 76.6667 | 73.8333 | 71 | 76 | 73.3333 | 79.3333 | 94 | 77.9833 | 73.6667 | 71.5302 | 63.3333 | 65.5 | 78.6667 | 51.3333 | 90.7028 | 63.8 | 48.8578 | 82.6667
| Median |  | 63.5087 | 78.8788 | 75.2 | 74.5 | 76.5 | 73.8 | 71.5 | 76 | 73 | 80 | 94 | 78.1 | 74.5 | 71.4393 | 60 | 65 | 77.5 | 51.5 | 91.2085 | 63.3 | 47.2071 | 83
| Std Dev. |  | 0.80732 | 0.70608 | 1.468 | 1.5166 | 2.3381 | 1.0152 | 1.2649 | 4.858 | 1.0328 | 2.0656 | 3.4641 | 0.93041 | 2.582 | 0.27195 | 5.164 | 1.2247 | 2.4221 | 2.4221 | 1.2541 | 6.0382 | 4.2058 | 1.0328
| Rel.Std.Dev. |  | 1.268% | 0.8957% | 1.959% | 2.036% | 3.05% | 1.375% | 1.782% | 6.392% | 1.408% | 2.604% | 3.685% | 1.193% | 3.505% | 0.3802% | 8.154% | 1.87% | 3.079% | 4.718% | 1.383% | 9.464% | 8.608% | 1.249%
| PDM3 |  | -13.81% | 6.707% | 1.45% | 0.8406% | 3.773% | -0.06178% | -3.897% | 2.871% | -0.7386% | 7.383% | 27.24% | 5.556% | -0.2874% | -3.179% | -14.27% | -11.34% | 6.48% | -30.52% | 22.77% | -13.64% | -33.87% | 11.89%
| Z-Score (Absolute) |  | 1.75 | 0.7 | 0.07 | 0 | 0.35 | 0.11 | 0.57 | 0.24 | 0.19 | 0.78 | 3.16 | 0.56 | 0.13 | 0.48 | NA | 1.46 | 0.67 | 3.75 | 2.62 | 1.73 | 4.15 | 1.32
Results from laboratory 15 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 926b (Certified Value 4.2 ppm) | [23 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 4.13 | 4.1 | 4.43 | 4.1 | 4.24 | 4.36 | 4.22 | 4.1 | 4.2 | 4 | 4.2 | 3.4544 | 4.3 | 3.67 | < 5 | 3.95 | 3.2174 | 3.9 |  |  |  |  | 1
| 1 | 2 | 4.21 | 4.08 | 4.37 | 4.02 | 4.19 | 4.47 | 4.3 | 4.2 | 4.1 | 4 | 4.2 | 3.6428 | 4.3 | 3.79 | < 5 | 3.85 | 3.1797 | 3.8 |  |  |  |  | #N/A
| 1 | 3 | 4.28 | 4.05 | 4.06 | 4.08 | 4.1 | 4.25 | 4.07 | 4.2 | 4.2 | 4.3 | 4.3 | 3.4871 | 4.4 | 3.83 | < 5 | 4.01 | 3.1479 | 4 |  |  |  |  | 16
| 1 | 4 | 4.31 | 4.11 | 4.02 | 4.07 | 4.23 | 4.46 | 4.08 | 4.1 | 4.4 | 4.2 | 4.3 | 3.448 | 4.3 | 3.96 | < 5 |  | 3.1997 | 3.8 |  |  |  |  | 4.2004
| 1 | 5 | 4.32 | 4.14 | 4.06 | 4.2 | 4.27 | 4.4 | 4.39 | 4.4 | 4.8 | 4.4 | 4.2 | 3.6829 | 4.3 | 3.88 | < 5 | 3.74 | 2.9072 | 3.9 |  |  |  |  | 88
| 1 | 6 | 4.33 | 4.18 | 3.94 | 4.17 | 4.37 | 4.45 | 4.16 | 4.3 | 4.4 | 4.2 | 4.4 | 3.5577 | 4.4 | 3.84 | < 5 | 3.61 | 3.1649 | 3.9
| Mean |  | 4.2633 | 4.11 | 4.1467 | 4.1067 | 4.2333 | 4.3983 | 4.2033 | 4.2167 | 4.35 | 4.1833 | 4.2667 | 3.5455 | 4.3333 | 3.8283 |  | 3.832 | 3.1361 | 3.8833
| Median |  | 4.295 | 4.105 | 4.06 | 4.09 | 4.235 | 4.425 | 4.19 | 4.2 | 4.3 | 4.2 | 4.25 | 3.5224 | 4.3 | 3.835 |  | 3.85 | 3.1723 | 3.9
| Std Dev. |  | 0.078401 | 0.045607 | 0.20196 | 0.066833 | 0.089144 | 0.083766 | 0.12596 | 0.1169 | 0.251 | 0.16021 | 0.08165 | 0.099707 | 0.05164 | 0.096626 |  | 0.161 | 0.11484 | 0.075277
| Rel.Std.Dev. |  | 1.839% | 1.11% | 4.87% | 1.627% | 2.106% | 1.904% | 2.997% | 2.772% | 5.77% | 3.83% | 1.914% | 2.812% | 1.192% | 2.524% |  | 4.201% | 3.662% | 1.938%
| PDM3 |  | 1.499% | -2.152% | -1.279% | -2.231% | 0.7844% | 4.713% | 0.0702% | 0.3876% | 3.562% | -0.4059% | 1.578% | -15.59% | 3.165% | -8.858% |  | -8.77% | -25.34% | -7.548%
| Z-Score (Absolute) |  | 0.49 | 0.27 | 0.09 | 0.29 | 0.34 | 1.17 | 0.19 | 0.26 | 0.92 | 0.09 | 0.51 | 3.09 | 0.84 | 1.68 | 8.32 | 1.66 | 5.14 | 1.41
| Analytical results for Cu in OREAS 926b (Certified Value 0.855 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.88626 | 0.88841 | 0.856 | 0.852 | 0.863 | 0.847 | 0.839 | 0.794 | 0.872 | 0.8757 | 0.86496 | 0.88402 | 0.834 | 0.84393 | 0.904 | 0.87414 | 0.8503 | 0.79674 | 0.82511 | 0.937 | 0.8421 | 0.868 | 1
| 1 | 2 | 0.88059 | 0.87183 | 0.855 | 0.857 | 0.884 | 0.841 | 0.827 | 0.831 | 0.883 | 0.8732 | 0.848 | 0.88259 | 0.826 | 0.84747 | 0.895 | 0.88216 | 0.8515 | 0.78914 | 0.81295 | 0.925 | 0.8379 | 0.86 | #N/A
| 1 | 3 | 0.89213 | 0.87425 | 0.856 | 0.839 | 0.845 | 0.852 | 0.804 | 0.848 | 0.844 | 0.8915 | 0.85006 | 0.89787 | 0.837 | 0.84636 | 0.891 | 0.86865 | 0.8649 | 0.8011 | 0.81133 | 0.926 | 0.8337 | 0.87 | 16
| 1 | 4 | 0.89196 | 0.87282 | 0.848 | 0.863 | 0.838 | 0.85 | 0.819 | 0.884 | 0.842 | 0.8898 | 0.87781 | 0.86717 | 0.842 | 0.84436 | 0.879 | 0.85737 | 0.8711 | 0.81038 | 0.82963 |  | 0.84105 | 0.842 | 0.85457
| 1 | 5 | 0.88481 | 0.88971 | 0.887 | 0.839 | 0.859 | 0.84 | 0.813 | 0.806 | 0.852 | 0.8589 | 0.87341 | 0.90301 | 0.833 | 0.84507 | 0.883 | 0.85465 | 0.8373 | 0.84655 | 0.8059 | 0.936 | 0.83055 | 0.859 | 89
| 1 | 6 | 0.89275 | 0.88212 | 0.874 | 0.867 | 0.818 | 0.842 | 0.836 | 0.804 | 0.851 | 0.8651 | 0.90622 | 0.88014 | 0.838 | 0.84808 | 0.89 | 0.87482 | 0.8479 | 0.85975 | 0.82127 | 0.937 | 0.86625 | 0.835
| Mean |  | 0.88808 | 0.87986 | 0.86267 | 0.85283 | 0.85117 | 0.84533 | 0.823 | 0.82783 | 0.85733 | 0.8757 | 0.87008 | 0.8858 | 0.835 | 0.84588 | 0.89033 | 0.86863 | 0.85383 | 0.81728 | 0.8177 | 0.9322 | 0.84192 | 0.85567
| Median |  | 0.88911 | 0.87819 | 0.856 | 0.8545 | 0.852 | 0.8445 | 0.823 | 0.8185 | 0.8515 | 0.87445 | 0.86919 | 0.88331 | 0.8355 | 0.84571 | 0.8905 | 0.87139 | 0.8509 | 0.80574 | 0.81711 | 0.936 | 0.83948 | 0.8595
| Std Dev. |  | 0.0049679 | 0.0080113 | 0.01472 | 0.011873 | 0.022763 | 0.0050465 | 0.01355 | 0.033944 | 0.016464 | 0.013033 | 0.021408 | 0.012919 | 0.0054406 | 0.0016951 | 0.0088468 | 0.010714 | 0.012224 | 0.028926 | 0.0090831 | 0.00614 | 0.012694 | 0.014152
| Rel.Std.Dev. |  | 0.5594% | 0.9105% | 1.706% | 1.392% | 2.674% | 0.597% | 1.646% | 4.1% | 1.92% | 1.488% | 2.46% | 1.458% | 0.6516% | 0.2004% | 0.9937% | 1.233% | 1.432% | 3.539% | 1.111% | 0.6587% | 1.508% | 1.654%
| PDM3 |  | 3.922% | 2.959% | 0.9477% | -0.203% | -0.398% | -1.081% | -3.694% | -3.128% | 0.3236% | 2.473% | 1.815% | 3.655% | -2.29% | -1.017% | 4.185% | 1.646% | -0.08597% | -4.364% | -4.314% | 9.084% | -1.479% | 0.1286%
| Z-Score (Absolute) |  | 1.28 | 0.97 | 0.3 | 0.07 | 0.14 | 0.36 | 1.22 | 1.03 | 0.1 | 0.81 | 0.59 | 1.19 | 0.76 | 0.34 | 1.37 | 0.53 | 0.04 | 1.44 | 1.42 | 2.98 | 0.49 | 0.04
| Analytical results for Dy in OREAS 926b (Certified Value 2.93 ppm) | [23 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
2
| 1 | 1 | 3.7 | 2.7 | 2.8 | 2.27 | 2.8182 | 3.2 | 2.59 | 3.16 | [14 empty] | 1
| 1 | 2 | 3.4 | 4 | 3 | 2.3 | 2.6086 | 3.4 | 2.84 | 3.16 | [14 empty] | #N/A
| 1 | 3 | 3.6 | 2.7 | 2.9 | 2.49 | 2.948 | 3.45 | 2.72 | 3.14 | [14 empty] | 16
| 1 | 4 | 3.4 | 2.3 | 3 | 2.4 | 2.4909 | 3.1 | 2.78 | 3.1 | [14 empty] | 2.9328
| 1 | 5 | 3.6 | 2.7 | 3.1 | 2.48 | 2.9456 | 3.2 | 2.84 | 3.24 | [14 empty] | 90
| 1 | 6 | 3.5 | 3.1 | 2.9 | 2.38 | 2.5436 | 3.2 | 2.73 | 3.15
| Mean |  | 3.5333 | 2.9167 | 2.95 | 2.3867 | 2.7258 | 3.2583 | 2.75 | 3.1583
| Median |  | 3.55 | 2.7 | 2.95 | 2.39 | 2.7134 | 3.2 | 2.755 | 3.155
| Std Dev. |  | 0.12111 | 0.58793 | 0.10488 | 0.090259 | 0.20413 | 0.13571 | 0.093808 | 0.045789
| Rel.Std.Dev. |  | 3.428% | 20.16% | 3.555% | 3.782% | 7.489% | 4.165% | 3.411% | 1.45%
| PDM3 |  | 20.48% | -0.5504% | 0.5861% | -18.62% | -7.058% | 11.1% | -6.233% | 7.69%
| Z-Score (Absolute) |  | 1.87 | 0.05 | 0.05 | 1.7 | 0.65 | 1.01 | 0.57 | 0.7
| Analytical results for Er in OREAS 926b (Certified Value 1.49 ppm) | [23 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
2
| 1 | 1 | 2 | 1.2 | 1.4 | 1.13 | 1.3191 | 1.75 | 1.22 | 1.74 | [14 empty] | 1
| 1 | 2 | 1.8 | 1.4 | 1.5 | 1.13 | 1.4698 | 1.8 | 1.39 | 1.69 | [14 empty] | 21
| 1 | 3 | 2 | 1.2 | 1.4 | 1.23 | 1.3064 | 1.95 | 1.39 | 1.74 | [14 empty] | 16
| 1 | 4 | 1.8 | 1.2 | 1.4 | 1.2 | 1.3507 | 1.6 | 1.35 | 1.7 | [14 empty] | 1.4932
| 1 | 5 | 1.9 | 1.2 | 1.7 | 1.25 | 1.4816 | 1.65 | 1.27 | 1.77 | [14 empty] | 91
| 1 | 6 | 1.8 | 1.1 | 1.4 | 1.18 | 1.3848 | 1.65 | 1.47 | 1.71
| Mean |  | 1.8833 | 1.2167 | 1.4667 | 1.1867 | 1.3854 | 1.7333 | 1.3483 | 1.725
| Median |  | 1.85 | 1.2 | 1.4 | 1.19 | 1.3677 | 1.7 | 1.37 | 1.725
| Std Dev. |  | 0.098319 | 0.098319 | 0.12111 | 0.050067 | 0.075116 | 0.1291 | 0.090425 | 0.030166
| Rel.Std.Dev. |  | 5.22% | 8.081% | 8.257% | 4.219% | 5.422% | 7.448% | 6.706% | 1.749%
| PDM3 |  | 26.13% | -18.52% | -1.775% | -20.53% | -7.218% | 16.08% | -9.7% | 15.53%
| Z-Score (Absolute) |  | 1.37 | 0.63 | 0.12 | 0.72 | 0.12 | 0.92 | 0.23 | 0.9
| Analytical results for Eu in OREAS 926b (Certified Value 1.11 ppm) | [23 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 17 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
3
| 1 | 1 | 1.22 | 1 | 1.1 | 1.06 | 1.15 | 1.02 | 1.1 | [15 empty] | 1
| 1 | 2 | 1.07 | 1 | 1 | 1.04 | 1.05 | 1.02 | 1.11 | [15 empty] | 22
| 1 | 3 | 1.07 | 0.9 | 1.1 | 1.17 | 1.05 | 0.99 | 1.11 | [15 empty] | 16
| 1 | 4 | 1.15 | 1 | 1.1 | 1.12 | 1.05 | 1.03 | 1.12 | [15 empty] | 1.111
| 1 | 5 | 1.11 | 1 | 1.2 | 1.17 | 1.15 | 1.14 | 1.15 | [15 empty] | 92
| 1 | 6 | 1.12 | 1 | 1.1 | 1.12 | 1.15 | 1.04 | 1.12
| Mean |  | 1.1233 | 0.98333 | 1.1 | 1.1133 | 1.1 | 1.04 | 1.1183
| Median |  | 1.115 | 1 | 1.1 | 1.12 | 1.1 | 1.025 | 1.115
| Std Dev. |  | 0.056451 | 0.040825 | 0.063246 | 0.054283 | 0.054772 | 0.051769 | 0.017224
| Rel.Std.Dev. |  | 5.025% | 4.152% | 5.75% | 4.876% | 4.979% | 4.978% | 1.54%
| PDM3 |  | 1.11% | -11.49% | -0.9901% | 0.21% | -0.9901% | -6.391% | 0.6601%
| Z-Score (Absolute) |  | 0.86 | 4.29 | 0 | 0.49 | 0 | 2.21 | 0.67
| Analytical results for Fe in OREAS 926b (Certified Value 3.69 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.8979 | 3.7429 | 3.59 | 3.58 | 3.63 | 3.59 | 3.52 | 3.77 | 3.81 | 3.73 | 3.67 | 3.6599 | 3.78 | 3.7068 | 3.62 | 3.78 | 3.66 | 3.4 | 3.8046 | 4 | 3.495 | 3.7 | 1
| 1 | 2 | 3.8396 | 3.6793 | 3.55 | 3.6 | 3.75 | 3.61 | 3.54 | 3.69 | 3.87 | 3.69 | 3.57 | 3.6142 | 3.9 | 3.7009 | 3.58 | 3.66 | 3.67 | 3.26 | 3.7793 | 3.89 | 3.478 | 3.679 | #N/A
| 1 | 3 | 3.8083 | 3.6761 | 3.59 | 3.5 | 3.48 | 3.62 | 3.49 | 3.66 | 3.78 | 3.79 | 3.59 | 3.688 | 3.88 | 3.7358 | 3.68 | 3.68 | 3.62 | 3.34 | 3.7547 | 3.98 | 3.391 | 3.707 | 16
| 1 | 4 | 3.884 | 3.6415 | 3.58 | 3.61 | 3.54 | 3.62 | 3.52 | 3.6 | 3.78 | 3.72 | 3.77 | 3.5747 | 3.82 | 3.7042 | 3.65 | 3.71 | 3.76 | 3.39 | 3.8333 |  | 3.411 | 3.672 | 3.686
| 1 | 5 | 3.925 | 3.7377 | 3.72 | 3.51 | 3.55 | 3.6 | 3.56 | 3.69 | 3.84 | 3.61 | 3.73 | 3.7162 | 3.87 | 3.7471 | 3.67 | 3.73 | 3.76 | 3.61 | 3.7235 | 4.06 | 3.365 | 3.707 | 93
| 1 | 6 | 3.8999 | 3.6914 | 3.64 | 3.65 | 3.44 | 3.58 | 3.63 | 3.75 | 3.83 | 3.65 | 3.9 | 3.6254 | 3.93 | 3.7424 | 3.66 | 3.77 | 3.88 | 3.65 | 3.8644 | 4.1 | 3.536 | 3.721
| Mean |  | 3.8758 | 3.6948 | 3.6117 | 3.575 | 3.565 | 3.6033 | 3.5433 | 3.6933 | 3.8183 | 3.6983 | 3.705 | 3.6464 | 3.8633 | 3.7229 | 3.6433 | 3.7217 | 3.725 | 3.4417 | 3.7933 | 4.006 | 3.446 | 3.6977
| Median |  | 3.891 | 3.6853 | 3.59 | 3.59 | 3.545 | 3.605 | 3.53 | 3.69 | 3.82 | 3.705 | 3.7 | 3.6426 | 3.875 | 3.7213 | 3.655 | 3.72 | 3.715 | 3.395 | 3.792 | 4 | 3.4445 | 3.7035
| Std Dev. |  | 0.043417 | 0.038963 | 0.06047 | 0.058907 | 0.11149 | 0.01633 | 0.048442 | 0.061536 | 0.035449 | 0.063377 | 0.12292 | 0.051776 | 0.05465 | 0.02109 | 0.037238 | 0.047924 | 0.094604 | 0.15459 | 0.051652 | 0.080498 | 0.066837 | 0.018608
| Rel.Std.Dev. |  | 1.12% | 1.055% | 1.674% | 1.648% | 3.127% | 0.4532% | 1.367% | 1.666% | 0.9284% | 1.714% | 3.318% | 1.42% | 1.415% | 0.5665% | 1.022% | 1.288% | 2.54% | 4.492% | 1.362% | 2.009% | 1.94% | 0.5032%
| PDM3 |  | 5.15% | 0.2397% | -2.016% | -3.011% | -3.282% | -2.242% | -3.87% | 0.1998% | 3.591% | 0.3355% | 0.5164% | -1.073% | 4.812% | 1.002% | -1.157% | 0.9685% | 1.059% | -6.628% | 2.912% | 8.682% | -6.51% | 0.3174%
| Z-Score (Absolute) |  | 1.36 | 0.01 | 0.64 | 0.92 | 1 | 0.71 | 1.16 | 0.02 | 0.93 | 0.02 | 0.07 | 0.38 | 1.27 | 0.2 | 0.4 | 0.19 | 0.22 | 1.93 | 0.74 | 2.35 | 1.9 | 0.01
| Analytical results for Ga in OREAS 926b (Certified Value 16.5 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 18.7671 | 16.9667 | 15 | 16.45 | 16.25 | 16.05 | 15.75 | 17.2 | 15.4 | 16.2 | 17.61 | 16.14 | 16 | 14.794 | 17.2 | 17 | 16.7 | 17.687 | 17.4 | 16.5268 | 15.9 |  | 1
| 1 | 2 | 17.3976 | 15.925 | 15.9 | 16.25 | 16.05 | 15.75 | 16 | 17 | 16.6 | 16.6 | 17.72 | 15.97 | 15.7 | 14.5298 | 16.6 | 16.9 | 16 | 17.137 | 16.6 | 16.7867 | 15.9 |  | #N/A
| 1 | 3 | 17.3735 | 15.9 | 16.3 | 16.3 | 15.55 | 16.05 | 15.75 | 17.2 | 15.9 | 16 | 17.33 | 16.2 | 16 | 14.8089 | 17.6 | 17 | 16.2 | 19.063 | 17.6 | 16.7634 | 16.4 |  | 16
| 1 | 4 | 18.5944 | 16.8083 | 15.75 | 16.3 | 15.2 | 15.9 | 16.25 | 16.9 | 16 | 16.6 | 18.09 | 15.93 | 16.2 | 14.7495 | 17.4 | 16.6 | 15.7 | 16.76 |  | 17.3382 | 15.8 |  | 16.53
| 1 | 5 | 16.4418 | 16.0917 | 16.5 | 16.6 | 15.35 | 15.7 | 16.55 | 16.9 | 16 | 14.7 | 17.1 | 16.5 | 16 | 14.9711 | 17.4 | 16.6 | 16.6 | 16.067 | 16.9 | 15.0179 | 16.2 |  | 94
| 1 | 6 | 18.7871 | 16.175 | 16.7 | 17 | 14.7 | 15.75 | 16.9 | 17 | 15.2 | 16.5 | 17.77 | 16.07 | 16.3 | 14.4817 | 17.8 | 16.6 | 17.5 | 16.874 | 16.9 | 16.6838 | 15.8
| Mean |  | 17.8936 | 16.3111 | 16.025 | 16.4833 | 15.5167 | 15.8667 | 16.2 | 17.0333 | 15.85 | 16.1 | 17.6033 | 16.135 | 16.0333 | 14.7225 | 17.3333 | 16.7833 | 16.45 | 17.2647 | 17.08 | 16.5194 | 16
| Median |  | 17.996 | 16.1333 | 16.1 | 16.375 | 15.45 | 15.825 | 16.125 | 17 | 15.95 | 16.35 | 17.665 | 16.105 | 16 | 14.7717 | 17.4 | 16.75 | 16.4 | 17.0055 | 16.9 | 16.7236 | 15.9
| Std Dev. |  | 0.96714 | 0.46082 | 0.61624 | 0.28402 | 0.56892 | 0.15706 | 0.46043 | 0.13663 | 0.49699 | 0.72664 | 0.34823 | 0.2054 | 0.20656 | 0.18452 | 0.41312 | 0.20412 | 0.63482 | 1.0267 | 0.40866 | 0.78526 | 0.24495
| Rel.Std.Dev. |  | 5.405% | 2.825% | 3.845% | 1.723% | 3.666% | 0.9899% | 2.842% | 0.8021% | 3.136% | 4.513% | 1.978% | 1.273% | 1.288% | 1.253% | 2.383% | 1.216% | 3.859% | 5.947% | 2.393% | 4.754% | 1.531%
| PDM3 |  | 8.249% | -1.324% | -3.055% | -0.2821% | -6.13% | -4.013% | -1.996% | 3.045% | -4.114% | -2.601% | 6.493% | -2.389% | -3.004% | -10.93% | 4.86% | 1.533% | -0.4838% | 4.445% | 3.327% | -0.06369% | -3.206%
| Z-Score (Absolute) |  | 2.4 | 0 | 0.43 | 0.26 | 1.21 | 0.67 | 0.17 | 1.1 | 0.7 | 0.32 | 1.96 | 0.27 | 0.42 | 2.41 | 1.55 | 0.72 | 0.21 | 1.45 | 1.17 | 0.32 | 0.47
| Analytical results for Gd in OREAS 926b (Certified Value 4.27 ppm) | [23 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 4.7 | 4 | 4.4 | 3.89 | 4.203 | 4.6 | 4.11 | 4.3 | [14 empty] | 1
| 1 | 2 | 4.6 | 3.9 | 4.4 | 3.87 | 4.1647 | 4.4 | 4.15 | 4.2 | [14 empty] | 6
| 1 | 3 | 4.6 | 3.7 | 4.4 | 4.23 | 4.2605 | 4.6 | 4.05 | 4.4 | [14 empty] | 16
| 1 | 4 | 4.7 | 3.9 | 4.5 | 4.1 | 4.4456 | 4.4 | 4.1 | 4.2 | [14 empty] | 4.2662
| 1 | 5 | 4.8 | 3.8 | 4.9 | 4.31 | 4.4171 | 4.4 | 4.06 | 4.3 | [14 empty] | 95
| 1 | 6 | 4.7 | 3.5 | 4.5 | 4.07 | 4.3263 | 4.4 | 4.08 | 4.2
| Mean |  | 4.6833 | 3.8 | 4.5167 | 4.0783 | 4.3029 | 4.4667 | 4.0917 | 4.2667
| Median |  | 4.7 | 3.85 | 4.45 | 4.085 | 4.2934 | 4.4 | 4.09 | 4.25
| Std Dev. |  | 0.075277 | 0.17889 | 0.19408 | 0.17668 | 0.11387 | 0.10328 | 0.03656 | 0.08165
| Rel.Std.Dev. |  | 1.607% | 4.708% | 4.297% | 4.332% | 2.646% | 2.312% | 0.8935% | 1.914%
| PDM3 |  | 9.778% | -10.93% | 5.871% | -4.403% | 0.8599% | 4.699% | -4.091% | 0.01111%
| Z-Score (Absolute) |  | 1.35 | 1.64 | 0.78 | 0.7 | 0.06 | 0.61 | 0.65 | 0.06
| Analytical results for Ge in OREAS 926b (Indicative Value 0.16 ppm) | [23 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 18 | 20 | 21 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 2
2
| 1 | 1 | 0.07 | 0.17 | 0.14 | 0.07 | 0.17 | < 0.1 | 1.14 | < 0.5 | 0.22 | 0.17 | 1.9 | 0.23 | < 0.1 | 0.24 | 0.27534 | [7 empty] | 1
| 1 | 2 | 0.07 | 0.17 | 0.12 | 0.07 | 0.15 | < 0.1 | 1.36 | < 0.5 | 0.25 | 0.17 | 1.9 | 0.22 | < 0.1 | 0.2 | 0.36472 | [7 empty] | 17
| 1 | 3 | 0.06 | 0.14 | 0.16 | 0.07 | 0.17 | < 0.1 | 1.21 | < 0.5 | 0.17 | 0.18 | 1.9 | 0.23 | < 0.1 | 0.23 | 0.35916 | [7 empty] | 16
| 1 | 4 | 0.05 | 0.17 | 0.13 | 0.07 | 0.16 | < 0.1 | 1.24 | < 0.5 | 0.31 | 0.16 | 1.8 | 0.24 | 0.2 |  | 0.31207 | [7 empty] | 0.16256
| 1 | 5 | 0.07 | 0.19 | 0.15 | 0.07 | 0.15 | < 0.1 | 1.24 | < 0.5 | 0.25 | 0.18 | 1.9 | 0.22 | 0.6 | 0.22 | 0.33604 | [7 empty] | 23
| 1 | 6 | 0.08 | 0.17 | 0.13 | 0.07 | 0.17 | < 0.1 | 1.18 | < 0.5 | 0.23 | 0.19 | 1.8 | 0.22 | 0.3 | 0.2 | 0.35772
| Mean |  | 0.066667 | 0.16833 | 0.13833 | 0.07 | 0.16167 |  | 1.2283 |  | 0.23833 | 0.175 | 1.8667 | 0.22667 | 0.36667 | 0.218 | 0.33417
| Median |  | 0.07 | 0.17 | 0.135 | 0.07 | 0.165 |  | 1.225 |  | 0.24 | 0.175 | 1.9 | 0.225 | 0.3 | 0.22 | 0.34688
| Std Dev. |  | 0.010328 | 0.016021 | 0.01472 | 0 | 0.0098319 |  | 0.074944 |  | 0.045789 | 0.010488 | 0.05164 | 0.008165 | 0.20817 | 0.017889 | 0.03483
| Rel.Std.Dev. |  | 15.49% | 9.517% | 10.64% | 0% | 6.082% |  | 6.101% |  | 19.21% | 5.993% | 2.766% | 3.602% | 56.77% | 8.206% | 10.42%
| PDM3 |  | -58.99% | 3.554% | -14.9% | -56.94% | -0.5468% |  | 655.6% |  | 46.62% | 7.656% | 1048% | 39.44% | 125.6% | 34.11% | 105.6%
| Z-Score (Absolute) |  | 1.15 | 0.07 | 0.39 | 1.12 | 0.14 | 1.33 | 11.21 | 0.8 | 0.67 | 0 | NA | 0.55 | NA | 0.46 | 1.69
Results from laboratories 13 and 18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 926b (Certified Value 2.48 ppm) | [23 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.5 | 2.6 | 2.5 | 2.37 | 2.6 | 2.9 | 2.4 | 2.5 | 3.03 | 1.89 | 2.5 | 2.2 | 2.4 | 2.8 | 2.44 | 2.4 | 2.3381 | 2.5 |  |  |  |  | 1
| 1 | 2 | 2.57 | 2.5 | 2.4 | 2.37 | 2.6 | 2.7 | 2.4 | 2.5 | 2.59 | 1.94 | 2.5 | 2.2 | 2.3 | 2.92 | 2.43 | 2.2 | 2.5047 | 2.4 |  |  |  |  | 13
| 1 | 3 | 2.46 | 2.5 | 2.6 | 2.25 | 2.6 | 3.1 | 2.2 | 2.5 | 2.42 | 2.1 | 2.4 | 2.2 | 2.5 | 2.73 | 2.49 | 2.4 | 2.189 | 2.6 |  |  |  |  | 16
| 1 | 4 | 2.56 | 2.5 | 2.5 | 2.31 | 2.7 | 2.7 | 2.4 | 2.4 | 2.6 | 2.02 | 2.4 | 2.4 | 2.3 | 2.7 | 2.36 |  | 2.1748 | 2.5 |  |  |  |  | 2.4768
| 1 | 5 | 2.53 | 2.5 | 2.4 | 2.44 | 2.5 | 2.7 | 2.3 | 2.3 | 2.66 | 2.09 | 2.4 | 2.4 | 2.5 | 2.78 | 2.52 | 2.4 | 2.1659 | 2.6 |  |  |  |  | 96
| 1 | 6 | 2.53 | 2.6 | 2.5 | 2.43 | 2.6 | 2.7 | 2.3 | 2.2 | 2.63 | 1.98 | 2.5 | 2.4 | 2.3 | 2.77 | 2.64 | 2.4 | 2.2814 | 2.5
| Mean |  | 2.525 | 2.5333 | 2.4833 | 2.3617 | 2.6 | 2.8 | 2.3333 | 2.4 | 2.655 | 2.0033 | 2.45 | 2.3 | 2.3833 | 2.7833 | 2.48 | 2.36 | 2.2757 | 2.5167
| Median |  | 2.53 | 2.5 | 2.5 | 2.37 | 2.6 | 2.7 | 2.35 | 2.45 | 2.615 | 2 | 2.45 | 2.3 | 2.35 | 2.775 | 2.465 | 2.4 | 2.2352 | 2.5
| Std Dev. |  | 0.040373 | 0.05164 | 0.075277 | 0.072226 | 0.063246 | 0.16733 | 0.08165 | 0.12649 | 0.20187 | 0.083106 | 0.054772 | 0.10954 | 0.098319 | 0.07607 | 0.095708 | 0.089443 | 0.13126 | 0.075277
| Rel.Std.Dev. |  | 1.599% | 2.038% | 3.031% | 3.058% | 2.433% | 5.976% | 3.499% | 5.27% | 7.603% | 4.148% | 2.236% | 4.763% | 4.125% | 2.733% | 3.859% | 3.79% | 5.768% | 2.991%
| PDM3 |  | 1.946% | 2.282% | 0.2636% | -4.649% | 4.974% | 13.05% | -5.793% | -3.101% | 7.195% | -19.12% | -1.082% | -7.138% | -3.774% | 12.38% | 0.1291% | -4.716% | -8.121% | 1.609%
| Z-Score (Absolute) |  | 0.39 | 0.44 | 0.12 | 0.67 | 0.87 | 2.17 | 0.85 | 0.42 | 1.23 | 2.99 | 0.1 | 1.07 | 0.53 | 2.06 | 0.1 | 0.68 | 1.23 | 0.33
| Analytical results for Hg in OREAS 926b (Indicative Value 0.032 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | Lab | [21 empty] | 1
| No | No | 12 | [21 empty] | ppm
|  |  | 4A*OES/MS | [21 empty] | 3
3
| 1 | 1 | 0.07 | [21 empty] | 1
| 1 | 2 | 0.03 | [21 empty] | 16
| 1 | 3 | 0.03 | [21 empty] | 16
| 1 | 4 | 0.03 | [21 empty] | 0.032
| 1 | 5 | 0.04 | [21 empty] | 22
| 1 | 6 | 0.03
| Mean |  | 0.038333
| Median |  | 0.03
| Std Dev. |  | 0.016021
| Rel.Std.Dev. |  | 41.79%
| PDM3 |  | 19.79%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 926b (Certified Value 0.54 ppm) | [23 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 17 | 22 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [15 empty] | 2
2
| 1 | 1 | 0.7 | 0.5 | 0.5 | 0.37 | 0.74 | 0.53 | 0.61 | [15 empty] | 1
| 1 | 2 | 0.7 | 0.5 | 0.5 | 0.37 | 0.62 | 0.52 | 0.6 | [15 empty] | 23
| 1 | 3 | 0.7 | 0.5 | 0.5 | 0.4 | 0.62 | 0.55 | 0.62 | [15 empty] | 16
| 1 | 4 | 0.7 | 0.4 | 0.5 | 0.39 | 0.58 | 0.56 | 0.6 | [15 empty] | 0.53762
| 1 | 5 | 0.6 | 0.4 | 0.5 | 0.4 | 0.58 | 0.54 | 0.62 | [15 empty] | 97
| 1 | 6 | 0.6 | 0.5 | 0.5 | 0.39 | 0.6 | 0.51 | 0.6
| Mean |  | 0.66667 | 0.46667 | 0.5 | 0.38667 | 0.62333 | 0.535 | 0.60833
| Median |  | 0.7 | 0.5 | 0.5 | 0.39 | 0.61 | 0.535 | 0.605
| Std Dev. |  | 0.05164 | 0.05164 | 0 | 0.013663 | 0.059889 | 0.018708 | 0.0098319
| Rel.Std.Dev. |  | 7.746% | 11.07% | 0% | 3.533% | 9.608% | 3.497% | 1.616%
| PDM3 |  | 24% | -13.2% | -6.997% | -28.08% | 15.94% | -0.4872% | 13.15%
| Z-Score (Absolute) |  | 1.21 | 0.63 | 0.32 | 1.36 | 0.81 | 0 | 0.67
| Analytical results for In in OREAS 926b (Certified Value 0.76 ppm) | [23 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 0.774 | 0.797 | 0.789 | 0.766 | 0.772 | 0.7 | 0.654 | 0.84 | 0.8 | 0.734 | 0.78 | 0.54685 | 0.75 | 0.73 | 0.72 | 0.716 | 0.74531 | [5 empty] | 1
| 1 | 2 | 0.766 | 0.774 | 0.768 | 0.735 | 0.749 | 0.7 | 0.721 | 0.74 | 0.75 | 0.78 | 0.78 | 0.52266 | 0.7 | 0.75 | 0.7 | 0.7 | 0.74118 | [5 empty] | #N/A
| 1 | 3 | 0.775 | 0.772 | 0.769 | 0.763 | 0.749 | 0.8 | 0.689 | 0.81 | 0.8 | 0.786 | 0.77 | 0.54492 | 0.75 | 0.75 | 0.71 | 0.714 | 0.73584 | [5 empty] | 16
| 1 | 4 | 0.796 | 0.821 | 0.76 | 0.773 | 0.735 | 0.7 | 0.663 | 0.92 | 0.84 | 0.735 | 0.77 | 0.55 | 0.75 | 0.76 | 0.69 |  | 0.72371 | [5 empty] | 0.75514
| 1 | 5 | 0.767 | 0.784 | 0.767 | 0.767 | 0.764 | 0.7 | 0.724 | 0.83 | 0.91 | 0.79 | 0.78 | 0.55918 | 0.75 | 0.76 | 0.75 | 0.75 | 0.65946 | [5 empty] | 98
| 1 | 6 | 0.762 | 0.822 | 0.74 | 0.767 | 0.756 | 0.7 | 0.697 | 0.78 | 0.86 | 0.815 | 0.79 | 0.54 | 0.75 | 0.75 | 0.77 | 0.741 | 0.73609
| Mean |  | 0.77333 | 0.795 | 0.7655 | 0.76183 | 0.75417 | 0.71667 | 0.69133 | 0.82 | 0.82667 | 0.77333 | 0.77833 | 0.54394 | 0.74167 | 0.75 | 0.72333 | 0.7242 | 0.7236
| Median |  | 0.7705 | 0.7905 | 0.7675 | 0.7665 | 0.7525 | 0.7 | 0.693 | 0.82 | 0.82 | 0.783 | 0.78 | 0.54589 | 0.75 | 0.75 | 0.715 | 0.716 | 0.73596
| Std Dev. |  | 0.01216 | 0.022361 | 0.015834 | 0.013541 | 0.012952 | 0.040825 | 0.028918 | 0.060992 | 0.055737 | 0.032358 | 0.0075277 | 0.012219 | 0.020412 | 0.010954 | 0.030768 | 0.020645 | 0.032251
| Rel.Std.Dev. |  | 1.572% | 2.813% | 2.068% | 1.777% | 1.717% | 5.696% | 4.183% | 7.438% | 6.742% | 4.184% | 0.9672% | 2.246% | 2.752% | 1.461% | 4.254% | 2.851% | 4.457%
| PDM3 |  | 2.409% | 5.278% | 1.371% | 0.8859% | -0.1294% | -5.095% | -8.45% | 8.589% | 9.471% | 2.409% | 3.071% | -27.97% | -1.785% | -0.6811% | -4.212% | -4.098% | -4.178%
| Z-Score (Absolute) |  | 0.43 | 0.92 | 0.26 | 0.17 | 0 | 0.84 | 1.41 | 1.48 | 1.63 | 0.43 | 0.54 | 4.73 | 0.28 | 0.09 | 0.69 | 0.67 | 0.69
| Analytical results for K in OREAS 926b (Certified Value 2.45 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.5244 | 2.4815 | 2.39 | 2.44 | 2.55 | 2.43 | 2.42 | 2.58 | 2.81 | 2.52 | 2.33 | 2.379 | 2.34 | 2.5222 | 2.49 | 2.4 | 2.51 | 2.33 | 2.71 | 2.43 | 2.222 | 2.457 | 1
| 1 | 2 | 2.5151 | 2.4439 | 2.39 | 2.43 | 2.49 | 2.45 | 2.46 | 2.89 | 2.78 | 2.5 | 2.31 | 2.4017 | 2.34 | 2.5146 | 2.51 | 2.33 | 2.46 | 2.33 | 2.7079 | 2.4 | 2.208 | 2.424 | #N/A
| 1 | 3 | 2.497 | 2.463 | 2.4 | 2.41 | 2.4 | 2.46 | 2.4 | 2.86 | 2.74 | 2.55 | 2.28 | 2.4576 | 2.34 | 2.5518 | 2.52 | 2.33 | 2.48 | 2.37 | 2.7118 | 2.44 | 2.188 | 2.457 | 16
| 1 | 4 | 2.5251 | 2.4299 | 2.38 | 2.47 | 2.47 | 2.46 | 2.44 | 2.86 | 2.73 | 2.51 | 2.39 | 2.373 | 2.38 | 2.5108 | 2.53 | 2.35 | 2.5 | 2.31 | 2.7835 |  | 2.186 | 2.424 | 2.4462
| 1 | 5 | 2.5255 | 2.4811 | 2.47 | 2.4 | 2.44 | 2.46 | 2.43 | 2.51 | 2.78 | 2.44 | 2.37 | 2.4369 | 2.37 | 2.5412 | 2.51 | 2.37 | 2.37 | 2.51 | 2.7714 | 2.41 | 2.16 | 2.449 | 99
| 1 | 6 | 2.5024 | 2.4603 | 2.45 | 2.48 | 2.34 | 2.42 | 2.47 | 1.94 | 2.75 | 2.48 | 2.41 | 2.4313 | 2.39 | 2.5236 | 2.55 | 2.4 | 2.51 | 2.61 | 2.7323 | 2.41 | 2.239 | 2.457
| Mean |  | 2.5149 | 2.46 | 2.4133 | 2.4383 | 2.4483 | 2.4467 | 2.4367 | 2.6067 | 2.765 | 2.5 | 2.3483 | 2.4133 | 2.36 | 2.5274 | 2.5183 | 2.3633 | 2.4717 | 2.41 | 2.7361 | 2.418 | 2.2005 | 2.4447
| Median |  | 2.5198 | 2.4616 | 2.395 | 2.435 | 2.455 | 2.455 | 2.435 | 2.72 | 2.765 | 2.505 | 2.35 | 2.4165 | 2.355 | 2.5229 | 2.515 | 2.36 | 2.49 | 2.35 | 2.722 | 2.41 | 2.198 | 2.453
| Std Dev. |  | 0.012504 | 0.020392 | 0.037238 | 0.031885 | 0.073052 | 0.017512 | 0.02582 | 0.36418 | 0.030166 | 0.037417 | 0.049967 | 0.033996 | 0.022804 | 0.015942 | 0.020412 | 0.032042 | 0.053448 | 0.12198 | 0.033412 | 0.016432 | 0.028311 | 0.016305
| Rel.Std.Dev. |  | 0.4972% | 0.8289% | 1.543% | 1.308% | 2.984% | 0.7157% | 1.06% | 13.97% | 1.091% | 1.497% | 2.128% | 1.409% | 0.9663% | 0.6308% | 0.8106% | 1.356% | 2.162% | 5.062% | 1.221% | 0.6796% | 1.287% | 0.667%
| PDM3 |  | 2.808% | 0.5613% | -1.345% | -0.3227% | 0.08611% | 0.01798% | -0.3908% | 6.559% | 13.03% | 2.198% | -4.002% | -1.348% | -3.525% | 3.318% | 2.948% | -3.389% | 1.04% | -1.481% | 11.85% | -1.154% | -10.05% | -0.06378%
| Z-Score (Absolute) |  | 1.04 | 0.21 | 0.49 | 0.11 | 0.04 | 0.01 | 0.14 | 2.41 | 4.78 | 0.81 | 1.46 | 0.49 | 1.28 | 1.22 | 1.09 | 1.23 | 0.39 | 0.54 | 4.35 | 0.42 | 3.67 | 0.01
| Analytical results for La in OREAS 926b (Certified Value 34.8 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  |  | 1
2
| 1 | 1 | 35.6 | 35.2 | 36 | 36.1 | 36.4 | 37.8 | 27.5 | 36.4 | 32.3 | 34.3 | 35.2 | 35.2 | 37.9 | 34 | 33.1 | 31.61 | 31.8 | 36 |  |  |  |  | 1
| 1 | 2 | 35.4 | 35.5 | 35 | 35.3 | 35.6 | 34.5 | 30.7 | 34.9 | 33.3 | 34.3 | 36.35 | 34.3 | 36.8 | 32 | 29.2 | 31.19 | 30.9 | 35 |  |  |  |  | #N/A
| 1 | 3 | 36 | 34.5 | 37 | 36 | 35.1 | 34.6 | 29.4 | 32.9 | 32.1 | 35 | 36.67 | 35.6 | 36.7 | 33 | 32.6 | 31.82 | 32.2 | 36 |  |  |  |  | 16
| 1 | 4 | 37 | 36.1 | 36.6 | 36.6 | 35.4 | 34.7 | 29.8 | 34.8 | 32.4 | 33.7 | 35.37 | 35.2 | 36.8 | 34 | 31.1 | 32.08 |  | 35 |  |  |  |  | 34.8033
| 1 | 5 | 37.2 | 34.8 | 36.8 | 36.2 | 36.1 | 34.8 | 30.5 | 32.9 | 35.1 | 34.9 | 36.14 | 34.6 | 38.6 | 33 | 33.8 | 32.42 | 34.8 | 36 |  |  |  |  | 100
| 1 | 6 | 36.4 | 36.1 | 35.6 | 37.5 | 36.4 | 34.9 | 30.2 | 34.2 | 32.6 | 34.5 | 34.78 | 35.9 | 36.6 | 34 | 37.2 | 32.27 | 33.4 | 35
| Mean |  | 36.2667 | 35.3667 | 36.1667 | 36.2833 | 35.8333 | 35.2167 | 29.6833 | 34.35 | 32.9667 | 34.45 | 35.7517 | 35.1333 | 37.2333 | 33.3333 | 32.8333 | 31.8983 | 32.62 | 35.5
| Median |  | 36.2 | 35.35 | 36.3 | 36.15 | 35.85 | 34.75 | 30 | 34.5 | 32.5 | 34.4 | 35.755 | 35.2 | 36.8 | 33.5 | 32.85 | 31.95 | 32.2 | 35.5
| Std Dev. |  | 0.73394 | 0.66232 | 0.77374 | 0.73052 | 0.5465 | 1.2734 | 1.1686 | 1.3368 | 1.1237 | 0.47223 | 0.74112 | 0.59889 | 0.82138 | 0.8165 | 2.6957 | 0.45481 | 1.5139 | 0.54772
| Rel.Std.Dev. |  | 2.024% | 1.873% | 2.139% | 2.013% | 1.525% | 3.616% | 3.937% | 3.892% | 3.409% | 1.371% | 2.073% | 1.705% | 2.206% | 2.449% | 8.21% | 1.426% | 4.641% | 1.543%
| PDM3 |  | 4.205% | 1.619% | 3.917% | 4.252% | 2.959% | 1.188% | -14.71% | -1.303% | -5.277% | -1.015% | 2.725% | 0.9482% | 6.982% | -4.224% | -5.66% | -8.347% | -6.273% | 2.002%
| Z-Score (Absolute) |  | 0.71 | 0.12 | 0.64 | 0.72 | 0.43 | 0.03 | 3.55 | 0.53 | 1.43 | 0.47 | 0.37 | 0.03 | 1.33 | 1.19 | 1.52 | 2.12 | 1.65 | 0.21
| Analytical results for Li in OREAS 926b (Certified Value 21.5 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 22 | 21.8 | 21.8 | 21.9 | 21.8 | 21 | 17.2 | 22 | 21.2 | 19 | 21.6 | 19.3942 | 20 | 21.7 | 21.09 | 27 | 27.39 | 23.5 | 21 |  |  |  | 1
| 1 | 2 | 21.7 | 21.8 | 22 | 22.2 | 22.2 | 21 | 17.3 | 22 | 20.6 | 19 | 22.3 | 19.3333 | 20 | 21 | 21.86 | 26 | 27.49 | 22.1 | 21 |  |  |  | #N/A
| 1 | 3 | 21.8 | 21.4 | 21.1 | 22.5 | 21.4 | 24 | 16.4 | 22 | 20.4 | 20 | 22.2 | 19.4611 | 20 | 21.3 | 22.12 | 27 | 27.49 | 23.8 | 21 |  |  |  | 16
| 1 | 4 | 21.8 | 22.3 | 21.3 | 22 | 21.7 | 22 | 16.4 | 22 | 20.9 | 19 | 22.3 | 19.3205 | 20 | 21.1 | 22.19 | 27 | 27.92 |  | 21 |  |  |  | 21.507
| 1 | 5 | 22.9 | 21.9 | 21.5 | 21.7 | 21.8 | 21 | 15.4 | 21 | 22 | 20 | 22.2 | 19.1781 | 30 | 21.2 | 22.69 | 29 | 28.07 | 20.2 | 21 |  |  |  | 101
| 1 | 6 | 21.9 | 22.3 | 20.8 | 21.9 | 22.6 | 21 | 17.2 | 21 | 20.2 | 19 | 21.8 | 19.1698 | 30 | 21.2 | 22.96 | 30 | 27.73 | 18.3 | 21
| Mean |  | 22.0167 | 21.9167 | 21.4167 | 22.0333 | 21.9167 | 21.6667 | 16.65 | 21.6667 | 20.8833 | 19.3333 | 22.0667 | 19.3095 | 23.3333 | 21.25 | 22.1517 | 27.6667 | 27.6817 | 21.58 | 21
| Median |  | 21.85 | 21.85 | 21.4 | 21.95 | 21.8 | 21 | 16.8 | 22 | 20.75 | 19 | 22.2 | 19.3269 | 20 | 21.2 | 22.155 | 27 | 27.61 | 22.1 | 21
| Std Dev. |  | 0.4446 | 0.34303 | 0.4446 | 0.28048 | 0.42151 | 1.2111 | 0.73689 | 0.5164 | 0.65243 | 0.5164 | 0.29439 | 0.11631 | 5.164 | 0.2429 | 0.65694 | 1.5055 | 0.27147 | 2.321 | 0
| Rel.Std.Dev. |  | 2.019% | 1.565% | 2.076% | 1.273% | 1.923% | 5.59% | 4.426% | 2.383% | 3.124% | 2.671% | 1.334% | 0.6023% | 22.13% | 1.143% | 2.966% | 5.442% | 0.9807% | 10.76% | 0%
| PDM3 |  | 2.37% | 1.905% | -0.4201% | 2.447% | 1.905% | 0.7423% | -22.58% | 0.7423% | -2.9% | -10.11% | 2.602% | -10.22% | 8.492% | -1.195% | 2.997% | 28.64% | 28.71% | 0.3393% | -2.357%
| Z-Score (Absolute) |  | 0.57 | 0.4 | 0.4 | 0.59 | 0.4 | 0 | 8.12 | 0 | 1.27 | 3.78 | 0.65 | 3.81 | 2.7 | 0.67 | 0.78 | 9.71 | 9.73 | 0.14 | 1.08
| Analytical results for Lu in OREAS 926b (Certified Value 0.21 ppm) | [23 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 17 | 18 | [15 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | [15 empty] | 2
3
| 1 | 1 | 0.3 | 0.22 | 0.2 | 0.19 | 0.24 | 0.21 | 0.21 | [15 empty] | 1
| 1 | 2 | 0.3 | 0.23 | 0.2 | 0.19 | 0.24 | 0.22 | 0.19 | [15 empty] | 24
| 1 | 3 | 0.3 | 0.22 | 0.2 | 0.21 | 0.24 | 0.23 | 0.21 | [15 empty] | 16
| 1 | 4 | 0.3 | 0.2 | 0.2 | 0.2 | 0.22 | 0.22 | 0.2 | [15 empty] | 0.21256
| 1 | 5 | 0.3 | 0.2 | 0.2 | 0.2 | 0.3 | 0.23 | 0.22 | [15 empty] | 102
| 1 | 6 | 0.3 | 0.2 | 0.2 | 0.2 | 0.22 | 0.23 | 0.23
| Mean |  | 0.3 | 0.21167 | 0.2 | 0.19833 | 0.24333 | 0.22333 | 0.21
| Median |  | 0.3 | 0.21 | 0.2 | 0.2 | 0.24 | 0.225 | 0.21
| Std Dev. |  | 0 | 0.013292 | 3.0405e-17 | 0.0075277 | 0.029439 | 0.008165 | 0.014142
| Rel.Std.Dev. |  | 0% | 6.279% | 1.52e-14% | 3.795% | 12.1% | 3.656% | 6.734%
| PDM3 |  | 41.14% | -0.4182% | -5.907% | -6.691% | 14.48% | 5.071% | -1.202%
| Z-Score (Absolute) |  | 5.11 | 0 | 0.67 | 0.77 | 1.83 | 0.67 | 0.1
| Analytical results for Mg in OREAS 926b (Certified Value 0.326 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.32669 | 0.31948 | 0.32 | 0.31 | 0.31 | 0.31 | 0.32 | 0.35 | 0.34 | 0.34 | 0.32 | 0.3401 | 0.328 | 0.35691 | 0.32 | 0.34 | 0.33 | 0.31 | 0.30488 | 0.33 | 0.33 | 0.356 | 1
| 1 | 2 | 0.32227 | 0.31281 | 0.32 | 0.3 | 0.32 | 0.31 | 0.32 | 0.33 | 0.35 | 0.34 | 0.31 | 0.3399 | 0.333 | 0.35266 | 0.32 | 0.34 | 0.31 | 0.29 | 0.30322 | 0.32 | 0.329 | 0.35 | #N/A
| 1 | 3 | 0.32022 | 0.31554 | 0.32 | 0.29 | 0.31 | 0.31 | 0.31 | 0.33 | 0.34 | 0.34 | 0.31 | 0.3458 | 0.334 | 0.35043 | 0.33 | 0.34 | 0.32 | 0.3 | 0.31041 | 0.33 | 0.323 | 0.362 | 16
| 1 | 4 | 0.32886 | 0.31357 | 0.32 | 0.3 | 0.31 | 0.31 | 0.31 | 0.33 | 0.34 | 0.34 | 0.33 | 0.3372 | 0.328 | 0.35628 | 0.32 | 0.33 | 0.32 | 0.3 | 0.31208 |  | 0.324 | 0.35 | 0.32634
| 1 | 5 | 0.32947 | 0.31947 | 0.33 | 0.3 | 0.31 | 0.31 | 0.32 | 0.33 | 0.34 | 0.33 | 0.32 | 0.4619 | 0.332 | 0.35715 | 0.33 | 0.34 | 0.32 | 0.33 | 0.33093 | 0.33 | 0.317 | 0.356 | 103
| 1 | 6 | 0.32046 | 0.31647 | 0.32 | 0.31 | 0.31 | 0.31 | 0.33 | 0.33 | 0.34 | 0.33 | 0.34 | 0.3383 | 0.332 | 0.35134 | 0.34 | 0.35 | 0.31 | 0.34 | 0.31132 | 0.33 | 0.331 | 0.35
| Mean |  | 0.32466 | 0.31622 | 0.32167 | 0.30167 | 0.31167 | 0.31 | 0.31833 | 0.33333 | 0.34167 | 0.33667 | 0.32167 | 0.36053 | 0.33117 | 0.35413 | 0.32667 | 0.34 | 0.31833 | 0.31167 | 0.31214 | 0.328 | 0.32567 | 0.354
| Median |  | 0.32448 | 0.31601 | 0.32 | 0.3 | 0.31 | 0.31 | 0.32 | 0.33 | 0.34 | 0.34 | 0.32 | 0.34 | 0.332 | 0.35447 | 0.325 | 0.34 | 0.32 | 0.305 | 0.31087 | 0.33 | 0.3265 | 0.353
| Std Dev. |  | 0.0041943 | 0.0028422 | 0.0040825 | 0.0075277 | 0.0040825 | 0 | 0.0075277 | 0.008165 | 0.0040825 | 0.005164 | 0.01169 | 0.049748 | 0.0025626 | 0.0030031 | 0.008165 | 0.0063246 | 0.0075277 | 0.019408 | 0.0098879 | 0.0044721 | 0.0053541 | 0.004899
| Rel.Std.Dev. |  | 1.292% | 0.8988% | 1.269% | 2.495% | 1.31% | 0% | 2.365% | 2.449% | 1.195% | 1.534% | 3.634% | 13.8% | 0.7738% | 0.848% | 2.499% | 1.86% | 2.365% | 6.227% | 3.168% | 1.363% | 1.644% | 1.384%
| PDM3 |  | -0.5126% | -3.099% | -1.431% | -7.559% | -4.495% | -5.006% | -2.452% | 2.144% | 4.698% | 3.166% | -1.431% | 10.48% | 1.481% | 8.516% | 0.1016% | 4.187% | -2.452% | -4.495% | -4.35% | 0.5101% | -0.2049% | 8.477%
| Z-Score (Absolute) |  | 0.03 | 0.59 | 0.23 | 1.55 | 0.89 | 1 | 0.45 | 0.54 | 1.09 | 0.76 | 0.23 | 2.33 | 0.4 | 1.91 | 0.1 | 0.98 | 0.45 | 0.89 | 0.86 | 0.19 | 0.03 | 1.89
| Analytical results for Mn in OREAS 926b (Certified Value 0.007 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 3
3
| 1 | 1 | 0.0078916 | 0.0071135 | 0.00728 | 0.0069 | 0.0073 | 0.00724 | 0.0073 | 0.0075 | 0.0077 | 0.0076 | 0.0077 | 0.0069 | 0.0075 | 0.006975 | 0.0076 | 0.008 | 0.0081 | < 0.01 | 0.0077503 | 0.00814 | 0.0074966 | 0.0085 | 1
| 1 | 2 | 0.0076562 | 0.006975 | 0.00714 | 0.007 | 0.0073 | 0.00718 | 0.0073 | 0.0078 | 0.0086 | 0.0075 | 0.0074 | 0.0067 | 0.0072 | 0.006978 | 0.0076 | 0.0076 | 0.0076 | < 0.01 | 0.007643 | 0.00785 | 0.00703 | 0.0077 | #N/A
| 1 | 3 | 0.0076562 | 0.0071031 | 0.00711 | 0.0068 | 0.007 | 0.0074 | 0.007 | 0.0076 | 0.009 | 0.0076 | 0.0074 | 0.0068 | 0.0073 | 0.006965 | 0.0076 | 0.0077 | 0.0074 | < 0.01 | 0.0076008 | 0.00788 | 0.0064439 | 0.0077 | 16
| 1 | 4 | 0.0078026 | 0.0071104 | 0.00701 | 0.007 | 0.0069 | 0.00726 | 0.0071 | 0.0074 | 0.0072 | 0.0075 | 0.0079 | 0.0066 | 0.0073 | 0.006986 | 0.0076 | 0.0077 | 0.0077 | < 0.01 | 0.0077589 |  | 0.0069149 | 0.007 | 0.0073895
| 1 | 5 | 0.0079296 | 0.007176 | 0.00731 | 0.0069 | 0.0072 | 0.00715 | 0.0071 | 0.0075 | 0.0072 | 0.0074 | 0.0077 | 0.0073 | 0.007 | 0.006938 | 0.0076 | 0.0079 | 0.0074 | < 0.01 | 0.0074785 | 0.00792 | 0.0069721 | 0.0077 | 104
| 1 | 6 | 0.0080324 | 0.0073531 | 0.00722 | 0.0071 | 0.0068 | 0.00717 | 0.0072 | 0.0076 | 0.0071 | 0.0074 | 0.0081 | 0.0066 | 0.0073 | 0.006961 | 0.008 | 0.0079 | 0.0074 | < 0.01 | 0.0076148 | 0.008 | 0.006805 | 0.007
| Mean |  | 0.0078281 | 0.0071385 | 0.0071783 | 0.00695 | 0.0070833 | 0.0072333 | 0.0071667 | 0.0075667 | 0.0078 | 0.0075 | 0.0077 | 0.0068167 | 0.0072667 | 0.0069672 | 0.0076667 | 0.0078 | 0.0076 |  | 0.007641 | 0.007958 | 0.0069438 | 0.0076
| Median |  | 0.0078471 | 0.007112 | 0.00718 | 0.00695 | 0.0071 | 0.00721 | 0.00715 | 0.00755 | 0.00745 | 0.0075 | 0.0077 | 0.00675 | 0.0073 | 0.00697 | 0.0076 | 0.0078 | 0.0075 |  | 0.0076289 | 0.00792 | 0.0069435 | 0.0077
| Std Dev. |  | 0.00015222 | 0.00012399 | 0.00011303 | 0.00010488 | 0.0002137 | 9.2014e-05 | 0.00012111 | 0.00013663 | 0.0008124 | 8.9443e-05 | 0.00027568 | 0.00026394 | 0.0001633 | 1.6893e-05 | 0.0001633 | 0.00015492 | 0.00027568 |  | 0.00010446 | 0.00011628 | 0.0003417 | 0.00055857
| Rel.Std.Dev. |  | 1.945% | 1.737% | 1.575% | 1.509% | 3.017% | 1.272% | 1.69% | 1.806% | 10.42% | 1.193% | 3.58% | 3.872% | 2.247% | 0.2425% | 2.13% | 1.986% | 3.627% |  | 1.367% | 1.461% | 4.921% | 7.35%
| PDM3 |  | 5.935% | -3.397% | -2.858% | -5.948% | -4.144% | -2.114% | -3.016% | 2.397% | 5.555% | 1.495% | 4.201% | -7.753% | -1.663% | -5.716% | 3.75% | 5.555% | 2.848% |  | 3.404% | 7.693% | -6.033% | 2.848%
| Z-Score (Absolute) |  | 1.03 | 0.57 | 0.47 | 1 | 0.69 | 0.35 | 0.5 | 0.42 | 0.96 | 0.27 | 0.73 | 1.31 | 0.27 | 0.96 | 0.66 | 0.96 | 0.5 | 5.51 | 0.6 | 1.33 | 1.02 | 0.54
| Analytical results for Mo in OREAS 926b (Certified Value 4.94 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 5.6491 | 4.6 | 4.71 | 5 | 5.08 | 4.77 | 4.88 | 5 | 5.58 | 5 | 5.14 | 4.5 | 4.8 | 6 | 5 | 6.26 | 5.12 | < 5 | 5.19 | 4.6458 | 5.52 |  | 1
| 1 | 2 | 6.2392 | 5.29 | 4.79 | 4.93 | 5.05 | 4.53 | 4.61 | 5 | 5.3 | 5 | 5.25 | 5.4 | 4.8 | 5.5 | 4.9 | 6.61 | 4.93 | < 5 | 4.83 | 4.8758 | 5.38 |  | 1
| 1 | 3 | 6.0976 | 4.8 | 4.9 | 4.88 | 4.88 | 4.54 | 4.57 | 5 | 5.02 | 5 | 4.84 | 4.4 | 4.9 | 5 | 4.8 | 6.31 | 4.99 | < 5 | 4.99 | 4.6475 | 5.4 |  | 16
| 1 | 4 | 6.3415 | 4.87 | 4.95 | 4.94 | 4.7 | 4.67 | 4.59 | 5 | 4.84 | 6 | 5.61 | 4.5 | 4.7 | 4.5 | 4.8 | 5.05 | 5 | < 5 |  | 4.7835 | 5.38 |  | 4.9405
| 1 | 5 | 5.3501 | 4.85 | 4.92 | 5.11 | 4.79 | 4.71 | 4.68 | 5 | 4.9 | 5 | 5.75 | 5.4 | 4.8 | 5 | 4.9 | 5.71 | 5.06 | < 5 | 5.27 | 4.5904 | 5.19 |  | 105
| 1 | 6 | 5.6884 | 4.83 | 4.95 | 5.07 | 4.65 | 4.73 | 4.77 | 5 | 4.84 | 5 | 5.49 | 4.1 | 4.8 | 5 | 5 | 5.53 | 5.42 | < 5 | 5.27 | 4.8677 | 5.18
| Mean |  | 5.8943 | 4.8733 | 4.87 | 4.9883 | 4.8583 | 4.6583 | 4.6833 | 5 | 5.08 | 5.1667 | 5.3467 | 4.7167 | 4.8 | 5.1667 | 4.9 | 5.9117 | 5.0867 |  | 5.11 | 4.7351 | 5.3417
| Median |  | 5.893 | 4.84 | 4.91 | 4.97 | 4.835 | 4.69 | 4.645 | 5 | 4.96 | 5 | 5.37 | 4.5 | 4.8 | 5 | 4.9 | 5.985 | 5.03 |  | 5.19 | 4.7155 | 5.38
| Std Dev. |  | 0.38958 | 0.22633 | 0.098184 | 0.088412 | 0.1786 | 0.10088 | 0.12094 | 0 | 0.29987 | 0.40825 | 0.33506 | 0.54924 | 0.063246 | 0.5164 | 0.089443 | 0.58249 | 0.17569 |  | 0.19391 | 0.12352 | 0.13212
| Rel.Std.Dev. |  | 6.609% | 4.644% | 2.016% | 1.772% | 3.676% | 2.166% | 2.582% | 0% | 5.903% | 7.902% | 6.267% | 11.64% | 1.318% | 9.995% | 1.825% | 9.853% | 3.454% |  | 3.795% | 2.609% | 2.473%
| PDM3 |  | 19.31% | -1.36% | -1.428% | 0.9676% | -1.664% | -5.712% | -5.206% | 1.204% | 2.823% | 4.577% | 8.221% | -4.531% | -2.844% | 4.577% | -0.8203% | 19.66% | 2.958% |  | 3.43% | -4.158% | 8.119%
| Z-Score (Absolute) |  | 3.59 | 0.1 | 0.11 | 0.32 | 0.15 | 0.87 | 0.78 | NA | 0.65 | NA | 1.61 | 0.66 | 0.36 | 0.96 | 0 | 3.65 | 0.67 | 8.67 | 0.76 | 0.6 | 1.6
Results from laboratories 8 and 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 926b (Certified Value 0.09 wt.%) | [23 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.10276 | 0.091775 | 0.087 | 0.09 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09 | 0.091 | 0.0778 | 0.085 | 0.11649 | 0.11 | 0.09 | 0.09 | 0.1 | 0.10122 | 0.1 | 0.09434 | 0.089 | 1
| 1 | 2 | 0.10067 | 0.089407 | 0.089 | 0.09 | 0.09 | 0.092 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.0778 | 0.084 | 0.11249 | 0.1 | 0.09 | 0.1 | 0.1 | 0.099658 | 0.1 | 0.09434 | 0.089 | 14
| 1 | 3 | 0.10284 | 0.090025 | 0.09 | 0.09 | 0.09 | 0.092 | 0.09 | 0.09 | 0.09 | 0.09 | 0.089 | 0.0796 | 0.086 | 0.11232 | 0.1 | 0.09 | 0.1 | 0.1 | 0.098377 | 0.1 | 0.09222 | 0.089 | 16
| 1 | 4 | 0.10446 | 0.089348 | 0.088 | 0.09 | 0.09 | 0.091 | 0.09 | 0.09 | 0.09 | 0.09 | 0.094 | 0.0765 | 0.084 | 0.11216 | 0.1 | 0.09 | 0.09 | 0.1 | 0.10204 |  | 0.09222 | 0.089 | 0.090217
| 1 | 5 | 0.1017 | 0.091602 | 0.092 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.091 | 0.0792 | 0.085 | 0.1136 | 0.11 | 0.09 | 0.08 | 0.11 | 0.099085 | 0.1 | 0.09116 | 0.089 | 106
| 1 | 6 | 0.10586 | 0.088642 | 0.091 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.096 | 0.078 | 0.087 | 0.11096 | 0.11 | 0.09 | 0.11 | 0.11 | 0.10066 | 0.1 | 0.09434 | 0.089
| Mean |  | 0.10305 | 0.090133 | 0.0895 | 0.09 | 0.09 | 0.090833 | 0.09 | 0.091667 | 0.09 | 0.09 | 0.091833 | 0.07815 | 0.085167 | 0.113 | 0.105 | 0.09 | 0.095 | 0.10333 | 0.10017 | 0.1 | 0.093103 | 0.089
| Median |  | 0.1028 | 0.089716 | 0.0895 | 0.09 | 0.09 | 0.0905 | 0.09 | 0.09 | 0.09 | 0.09 | 0.091 | 0.0779 | 0.085 | 0.11241 | 0.105 | 0.09 | 0.095 | 0.1 | 0.10016 | 0.1 | 0.09328 | 0.089
| Std Dev. |  | 0.001873 | 0.0012831 | 0.0018708 | 1.5202e-17 | 1.5202e-17 | 0.00098319 | 1.5202e-17 | 0.0040825 | 1.5202e-17 | 1.5202e-17 | 0.0026394 | 0.0011131 | 0.001169 | 0.0019027 | 0.0054772 | 1.5202e-17 | 0.010488 | 0.005164 | 0.0013776 | 0 | 0.0014089 | 1.5202e-17
| Rel.Std.Dev. |  | 1.818% | 1.424% | 2.09% | 1.689e-14% | 1.689e-14% | 1.082% | 1.689e-14% | 4.454% | 1.689e-14% | 1.689e-14% | 2.874% | 1.424% | 1.373% | 1.684% | 5.216% | 1.689e-14% | 11.04% | 4.997% | 1.375% | 0% | 1.513% | 1.708e-14%
| PDM3 |  | 14.23% | -0.09309% | -0.7949% | -0.2407% | -0.2407% | 0.683% | -0.2407% | 1.607% | -0.2407% | -0.2407% | 1.791% | -13.38% | -5.598% | 25.26% | 16.39% | -0.2407% | 5.301% | 14.54% | 11.04% | 10.84% | 3.199% | -1.349%
| Z-Score (Absolute) |  | 5.68 | 0 | 0.28 | 0.06 | 0.06 | 0.31 | 0.06 | 0.67 | 0.06 | 0.06 | 0.75 | 5.27 | 2.18 | 10.06 | 6.54 | 0.06 | 2.14 | 5.8 | 4.42 | NA | 1.31 | 0.49
Results from laboratory 20 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb in OREAS 926b (Certified Value 4.72 ppm) | [23 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 5.28 | 5.1 | 4.8 | 5.07 | 4.8 | 8.4 | 3.7 | 3.3 | 5.42 | 2.64 | 4.6 | 4.5865 | 8 | 4.3 | 6.37 | 6.1 | 3.5 | 4.0177 | 10.3 |  |  |  | 1
| 1 | 2 | 5.27 | 4.8 | 4.4 | 4.76 | 5.2 | 8.8 | 4.2 | 3 | 4.75 | 2.8 | 4.4 | 4.5476 | 8 | 4.2 | 6.36 | 6.1 | 3.3 | 3.9933 | 10 |  |  |  | 3
| 1 | 3 | 5.48 | 4.8 | 4.5 | 4.84 | 5.5 | 10.1 | 3.8 | 2.8 | 4.55 | 2.9 | 4 | 4.5864 | 8.5 | 4.1 | 6.25 | 6.6 | 3.6 | 3.8894 | 10.1 |  |  |  | 16
| 1 | 4 | 5.52 | 5.7 | 4.6 | 4.95 | 5.5 | 7.6 | 4.3 | 3.3 | 4.95 | 2.69 | 4.2 | 4.564 | 9 | 4.2 | 6.17 | 5.5 |  | 3.7574 | 9.9 |  |  |  | 4.7212
| 1 | 5 | 5.65 | 5 | 4.4 | 5.07 | 5.5 | 9.6 | 3.8 | 3 | 4.78 | 2.92 | 4.1 | 4.5932 | 9.5 | 4.3 | 6.97 | 5.9 | 4.3 | 3.6928 | 10.3 |  |  |  | 107
| 1 | 6 | 5.35 | 5.2 | 4.6 | 5.18 | 6.3 | 10.6 | 3.8 | 3.2 | 4.99 | 2.98 | 4.3 | 4.5209 | 9.5 | 4.3 | 6.8 | 6.5 | 4.2 | 4.1011 | 10
| Mean |  | 5.425 | 5.1 | 4.55 | 4.9783 | 5.4667 | 9.1833 | 3.9333 | 3.1 | 4.9067 | 2.8217 | 4.2667 | 4.5664 | 8.75 | 4.2333 | 6.4867 | 6.1167 | 3.78 | 3.9086 | 10.1
| Median |  | 5.415 | 5.05 | 4.55 | 5.01 | 5.5 | 9.2 | 3.8 | 3.1 | 4.865 | 2.85 | 4.25 | 4.5752 | 8.75 | 4.25 | 6.365 | 6.1 | 3.6 | 3.9413 | 10.05
| Std Dev. |  | 0.15057 | 0.33466 | 0.15166 | 0.15817 | 0.49261 | 1.1215 | 0.25033 | 0.2 | 0.29669 | 0.13541 | 0.21602 | 0.028118 | 0.6892 | 0.08165 | 0.32179 | 0.40208 | 0.44385 | 0.1587 | 0.16733
| Rel.Std.Dev. |  | 2.775% | 6.562% | 3.333% | 3.177% | 9.011% | 12.21% | 6.364% | 6.452% | 6.047% | 4.799% | 5.063% | 0.6158% | 7.877% | 1.929% | 4.961% | 6.573% | 11.74% | 4.06% | 1.657%
| PDM3 |  | 14.91% | 8.023% | -3.627% | 5.446% | 15.79% | 94.51% | -16.69% | -34.34% | 3.928% | -40.23% | -9.628% | -3.279% | 85.33% | -10.33% | 37.39% | 29.56% | -19.94% | -17.21% | 113.9%
| Z-Score (Absolute) |  | 0.36 | 0.13 | 0.25 | 0.05 | 0.39 | 2.96 | 0.67 | 1.25 | 0 | 1.44 | 0.44 | 0.24 | 2.66 | 0.47 | 1.09 | 0.84 | 0.78 | 0.69 | 3.6
| Analytical results for Nd in OREAS 926b (Certified Value 29.1 ppm) | [23 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 1
2
| 1 | 1 | 30.7 | 29.2 | 32 | 26.33 | 29.2789 | 30 | 28.3 | 28 | [14 empty] | 1
| 1 | 2 | 28.6 | 29.2 | 31.8 | 26.22 | 29.7686 | 29 | 29.3 | 27.6 | [14 empty] | 4
| 1 | 3 | 31.5 | 26.7 | 31.6 | 28.57 | 29.815 | 29.9 | 29.2 | 28.9 | [14 empty] | 16
| 1 | 4 | 29.6 | 28.9 | 32.4 | 27.77 | 29.5201 | 29.5 | 28.5 | 27.4 | [14 empty] | 29.1395
| 1 | 5 | 29.3 | 26.6 | 35.1 | 28.92 | 29.1495 | 29.2 | 27.2 | 28.1 | [14 empty] | 108
| 1 | 6 | 30.1 | 27.8 | 32.5 | 27.76 | 29.3318 | 30.3 | 27.6 | 27.7
| Mean |  | 29.9667 | 28.0667 | 32.5667 | 27.595 | 29.4773 | 29.65 | 28.35 | 27.95
| Median |  | 29.85 | 28.35 | 32.2 | 27.765 | 29.4259 | 29.7 | 28.4 | 27.85
| Std Dev. |  | 1.0347 | 1.2127 | 1.2879 | 1.1185 | 0.27157 | 0.501 | 0.84083 | 0.53198
| Rel.Std.Dev. |  | 3.453% | 4.321% | 3.955% | 4.053% | 0.9213% | 1.69% | 2.966% | 1.903%
| PDM3 |  | 2.839% | -3.682% | 11.76% | -5.3% | 1.159% | 1.752% | -2.709% | -4.082%
| Z-Score (Absolute) |  | 0.78 | 0.63 | 2.72 | 0.98 | 0.42 | 0.55 | 0.42 | 0.72
| Analytical results for Ni in OREAS 926b (Certified Value 8.84 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 12.6148 | 8.07 | 8.43 | 8.8 | 9.1 | 8.42 | 8.8 | 8 | 11.6 | < 10 | 9.6 | 8.4 | 9 | 9.45 | 8 | 9.5 | 8.84 | 6.6 | 9.127 | 9.5 | 11.5818 | 10.1 | 1
| 1 | 2 | 12.4201 | 8.2 | 8.68 | 8.6 | 9.6 | 8.39 | 9 | 9 | 11.1 | < 10 | 10.1 | 8.2 | 9 | 9.28 | 8 | 9.3 | 8.54 | 7.1 | 8.314 | 8.4 | 14.5904 | 9.9 | 5
| 1 | 3 | 13.653 | 8.14 | 8.76 | 8.4 | 8.7 | 8.45 | 8.7 | 9 | 11.8 | < 10 | 10.7 | 8.5 | 9 | 9.74 | 8 | 9.3 | 8.69 | 7.3 | 8.892 | 8.8 | 11.8647 | 10.1 | 16
| 1 | 4 | 13.5492 | 7.83 | 8.95 | 8.5 | 8.1 | 8.62 | 8.6 | 9 | 9.9 | < 10 | 10.9 | 8.3 | 9 | 9.41 | 8 | 9.4 | 8.58 | 7.2 | 9.336 |  | 11.9539 | 10 | 8.8378
| 1 | 5 | 12.5239 | 8.45 | 8.95 | 8.5 | 8.3 | 8.62 | 8.6 | 8 | 10.9 | < 10 | 10.8 | 8.7 | 9 | 9.54 | 8 | 9.4 | 8.48 | 7.5 | 8.482 | 9 | 11.0876 | 10.1 | 109
| 1 | 6 | 12.5446 | 8.37 | 9.15 | 8.6 | 8 | 8.7 | 9 | 8 | 10.7 | < 10 | 9.6 | 8.4 | 9 | 9.31 | 8 | 9.4 | 8.15 | 7.1 | 9.013 | 12 | 11.8969 | 10.1
| Mean |  | 12.8843 | 8.1767 | 8.82 | 8.5667 | 8.6333 | 8.5333 | 8.7833 | 8.5 | 11 |  | 10.2833 | 8.4167 | 9 | 9.455 | 8 | 9.3833 | 8.5467 | 7.1333 | 8.8607 | 9.54 | 12.1626 | 10.05
| Median |  | 12.5797 | 8.17 | 8.855 | 8.55 | 8.5 | 8.535 | 8.75 | 8.5 | 11 |  | 10.4 | 8.4 | 9 | 9.43 | 8 | 9.4 | 8.56 | 7.15 | 8.9525 | 9 | 11.8808 | 10.1
| Std Dev. |  | 0.55971 | 0.2216 | 0.25219 | 0.13663 | 0.62503 | 0.12894 | 0.18348 | 0.54772 | 0.68118 |  | 0.59805 | 0.17224 | 0 | 0.16861 | 0 | 0.075277 | 0.23218 | 0.30111 | 0.39072 | 1.4311 | 1.2321 | 0.083666
| Rel.Std.Dev. |  | 4.344% | 2.71% | 2.859% | 1.595% | 7.24% | 1.511% | 2.089% | 6.444% | 6.193% |  | 5.816% | 2.046% | 0% | 1.783% | 0% | 0.8022% | 2.717% | 4.221% | 4.41% | 15% | 10.13% | 0.8325%
| PDM3 |  | 45.79% | -7.481% | -0.2017% | -3.068% | -2.314% | -3.445% | -0.6165% | -3.822% | 24.47% |  | 16.36% | -4.765% | 1.835% | 6.983% | -9.48% | 6.172% | -3.294% | -19.29% | 0.2585% | 7.945% | 37.62% | 13.72%
| Z-Score (Absolute) |  | 4.32 | 0.68 | 0 | 0.27 | 0.2 | 0.3 | 0.04 | NA | 2.31 | 4.06 | 1.55 | 0.43 | NA | 0.67 | NA | 0.6 | 0.29 | 1.79 | 0.04 | 0.76 | 3.55 | 1.31
Results from laboratories 8, 13 and 15 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 926b (Certified Value 0.025 wt.%) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.026019 | 0.023898 | 0.025 | 0.025 | 0.025 | 0.024 | 0.025 | 0.024 | 0.02 | 0.03 | 0.025 | 0.0253 | 0.025 | 0.049541 | 0.025 | 0.02 | 0.0259 | 0.026 | 0.02439 | 0.0272 | 0.026 | 0.0262 | 1
| 1 | 2 | 0.025631 | 0.023757 | 0.025 | 0.025 | 0.026 | 0.023 | 0.025 | 0.027 | 0.019 | 0.03 | 0.024 | 0.0245 | 0.025 | 0.049799 | 0.025 | 0.02 | 0.0275 | 0.0255 | 0.024782 | 0.0264 | 0.028 | 0.0262 | #N/A
| 1 | 3 | 0.024659 | 0.024221 | 0.025 | 0.025 | 0.026 | 0.024 | 0.025 | 0.027 | 0.02 | 0.03 | 0.024 | 0.0257 | 0.025 | 0.04975 | 0.03 | 0.02 | 0.0277 | 0.0244 | 0.024282 | 0.0273 | 0.026 | 0.0262 | 16
| 1 | 4 | 0.026201 | 0.02386 | 0.025 | 0.026 | 0.025 | 0.023 | 0.024 | 0.028 | 0.021 | 0.03 | 0.026 | 0.025 | 0.025 | 0.049477 | 0.03 | 0.02 | 0.026 | 0.0252 | 0.025913 |  | 0.027 | 0.0262 | 0.025417
| 1 | 5 | 0.026047 | 0.024014 | 0.026 | 0.025 | 0.025 | 0.023 | 0.025 | 0.024 | 0.018 | 0.03 | 0.025 | 0.0264 | 0.025 | 0.049348 | 0.025 | 0.02 | 0.0286 | 0.0272 | 0.025008 | 0.026 | 0.025 | 0.0262 | 110
| 1 | 6 | 0.026177 | 0.024428 | 0.025 | 0.026 | 0.025 | 0.023 | 0.026 | 0.025 | 0.02 | 0.03 | 0.027 | 0.0256 | 0.025 | 0.049288 | 0.03 | 0.02 | 0.0264 | 0.027 | 0.025204 | 0.0269 | 0.026 | 0.0218
| Mean |  | 0.025789 | 0.02403 | 0.025167 | 0.025333 | 0.025333 | 0.023333 | 0.025 | 0.025833 | 0.019667 | 0.03 | 0.025167 | 0.025417 | 0.025 | 0.049534 | 0.0275 | 0.02 | 0.027017 | 0.025883 | 0.02493 | 0.02676 | 0.026333 | 0.025467
| Median |  | 0.026033 | 0.023956 | 0.025 | 0.025 | 0.025 | 0.023 | 0.025 | 0.026 | 0.02 | 0.03 | 0.025 | 0.02545 | 0.025 | 0.049509 | 0.0275 | 0.02 | 0.02695 | 0.02575 | 0.024895 | 0.0269 | 0.026 | 0.0262
| Std Dev. |  | 0.00059027 | 0.00025139 | 0.00040825 | 0.0005164 | 0.0005164 | 0.0005164 | 0.00063246 | 0.0017224 | 0.0010328 | 0 | 0.001169 | 0.00064936 | 3.8006e-18 | 0.00020771 | 0.0027386 | 0 | 0.0010834 | 0.0010778 | 0.00059655 | 0.00055045 | 0.0010328 | 0.0017963
| Rel.Std.Dev. |  | 2.289% | 1.046% | 1.622% | 2.038% | 2.038% | 2.213% | 2.53% | 6.667% | 5.252% | 0% | 4.645% | 2.555% | 1.52e-14% | 0.4193% | 9.959% | 0% | 4.01% | 4.164% | 2.393% | 2.057% | 3.922% | 7.054%
| PDM3 |  | 1.463% | -5.459% | -0.9853% | -0.3296% | -0.3296% | -8.198% | -1.641% | 1.638% | -22.62% | 18.03% | -0.9853% | -0.001726% | -1.641% | 94.88% | 8.195% | -21.31% | 6.293% | 1.834% | -1.917% | 5.283% | 3.605% | 0.195%
| Z-Score (Absolute) |  | 0.58 | 1.88 | 0.29 | 0.06 | 0.06 | 2.85 | 0.52 | 0.64 | 7.96 | 6.45 | 0.29 | 0.06 | 0.52 | 33.7 | 2.96 | 7.5 | 2.29 | 0.71 | 0.62 | 1.93 | 1.34 | 0.11
| Analytical results for Pb in OREAS 926b (Certified Value 51 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 59.34 | 42.13 | 58.9 | 48.6 | 46.1 | 52.4 | 50 | 46 | 184 | 70 | 50.75 | 98 | 51 | 56.75 | 44 | 55.3 | 54 | 65.9 | 62.96 | 63.9 | 51.4377 | 235 | 1
| 1 | 2 | 53.98 | 43.88 | 45 | 47.4 | 46.3 | 48.3 | 45.2 | 47 | 98.8 | 59 | 49.13 | 61 | 53 | 55.2925 | 40 | 51.4 | 57 | 42.7 | 52.76 | 49 | 50.2293 | 81 | 6
| 1 | 3 | 54 | 45.66 | 46.1 | 47.2 | 49.8 | 47.8 | 44.5 | 43 | 135 | 56 | 48.7 | 62 | 54 | 53.413 | 58 | 50.9 | 52 | 42.3 | 48.52 | 49.4 | 47.5621 | 48.6 | 16
| 1 | 4 | 55.33 | 44.18 | 44.9 | 47.7 | 46.5 | 47 | 45.2 | 46 | 73.3 | 52 | 50.75 | 57 | 53 | 54.1835 | 54 | 50.6 | 53 | 39.1 | 47.15 |  | 48.3782 | 45.9 | 51.1048
| 1 | 5 | 59.33 | 42.07 | 46.4 | 47.4 | 46.3 | 47.6 | 44.7 | 45 | 77 | 49 | 50.35 | 60 | 52 | 57.537 | 44 | 51 | 51 | 40.6 | 50.25 | 46.2 | 46.3214 | 66.8 | 111
| 1 | 6 | 63.44 | 47.21 | 45.8 | 48.5 | 45.1 | 47.9 | 44.8 | 46 | 98.4 | 50 | 48.44 | 59 | 51 | 56.361 | 52 | 50.8 | 54 | 45.4 | 55.45 | 45.6 | 46.3439 | 46.3
| Mean |  | 57.57 | 44.1883 | 47.85 | 47.8 | 46.6833 | 48.5 | 45.7333 | 45.5 | 111.0833 | 56 | 49.6867 | 66.1667 | 52.3333 | 55.5895 | 48.6667 | 51.6667 | 53.5 | 46 | 52.8483 | 50.82 | 48.3788 | 87.2667
| Median |  | 57.33 | 44.03 | 45.95 | 47.55 | 46.3 | 47.85 | 45 | 46 | 98.6 | 54 | 49.74 | 60.5 | 52.5 | 55.8267 | 48 | 50.95 | 53.5 | 42.5 | 51.505 | 49 | 47.9701 | 57.7
| Std Dev. |  | 3.7772 | 2.0058 | 5.446 | 0.60332 | 1.6055 | 1.9575 | 2.1087 | 1.3784 | 41.9281 | 7.823 | 1.0525 | 15.6897 | 1.2111 | 1.5835 | 7.0048 | 1.7996 | 2.0736 | 9.9772 | 5.7789 | 7.5001 | 2.0891 | 73.7128
| Rel.Std.Dev. |  | 6.561% | 4.539% | 11.38% | 1.262% | 3.439% | 4.036% | 4.611% | 3.029% | 37.74% | 13.97% | 2.118% | 23.71% | 2.314% | 2.849% | 14.39% | 3.483% | 3.876% | 21.69% | 10.93% | 14.76% | 4.318% | 84.47%
| PDM3 |  | 12.65% | -13.53% | -6.369% | -6.467% | -8.652% | -5.097% | -10.51% | -10.97% | 117.4% | 9.579% | -2.775% | 29.47% | 2.404% | 8.775% | -4.771% | 1.099% | 4.687% | -9.989% | 3.412% | -0.5574% | -5.334% | 70.76%
| Z-Score (Absolute) |  | 1.45 | 1.2 | 0.48 | 0.49 | 0.71 | 0.35 | 0.89 | 0.94 | 12.03 | 1.14 | 0.11 | 3.15 | 0.41 | 1.06 | 0.31 | 0.28 | 0.64 | 0.84 | 0.51 | 0.11 | 0.37 | 7.32
| Analytical results for Pr in OREAS 926b (Certified Value 7.95 ppm) | [23 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 8.9 | 8.28 | 8.1 | 7.16 | 6.1817 | 8.2 | 8.13 | 7.7 | [14 empty] | 1
| 1 | 2 | 8.3 | 7.57 | 8.3 | 7.16 | 6.2683 | 7.95 | 8.15 | 7.6 | [14 empty] | 7
| 1 | 3 | 8.5 | 7.29 | 8.1 | 7.78 | 6.2881 | 8.25 | 8.4 | 7.9 | [14 empty] | 16
| 1 | 4 | 8.4 | 7.39 | 8.4 | 7.55 | 6.2886 | 8.15 | 8.25 | 7.5 | [14 empty] | 7.9523
| 1 | 5 | 8.1 | 7.12 | 8.9 | 7.84 | 6.6533 | 8.1 | 8.23 | 7.6 | [14 empty] | 112
| 1 | 6 | 8.4 | 7.51 | 8.4 | 7.58 | 6.1607 | 8.25 | 8.01 | 7.5
| Mean |  | 8.4333 | 7.5267 | 8.3667 | 7.5117 | 6.3068 | 8.15 | 8.195 | 7.6333
| Median |  | 8.4 | 7.45 | 8.35 | 7.565 | 6.2782 | 8.175 | 8.19 | 7.6
| Std Dev. |  | 0.26583 | 0.40243 | 0.29439 | 0.29438 | 0.17844 | 0.11402 | 0.13172 | 0.15055
| Rel.Std.Dev. |  | 3.152% | 5.347% | 3.519% | 3.919% | 2.829% | 1.399% | 1.607% | 1.972%
| PDM3 |  | 6.049% | -5.352% | 5.211% | -5.541% | -20.69% | 2.486% | 3.052% | -4.011%
| Z-Score (Absolute) |  | 0.98 | 0.66 | 0.86 | 0.69 | 2.87 | 0.47 | 0.55 | 0.47
| Analytical results for Rb in OREAS 926b (Certified Value 142 ppm) | [23 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 148 | 142 | 149.5 | 144.5 | 139 | 151 | 130 | 146.8 | 149.2 | 132.9 | 131 | 118.1135 | 144 | 147.5 | 145 | 126.6 | 135 | 134.5862 | 141 |  |  |  | 1
| 1 | 2 | 149.5 | 140.5 | 148.5 | 140 | 138.5 | 164 | 136 | 143.1 | 151.6 | 139.4 | 131 | 118.2179 | 140 | 144.6 | 144 | 124.8 | 130 | 141.1366 | 138 |  |  |  | 8
| 1 | 3 | 151.5 | 139.5 | 142.5 | 142.5 | 136.5 | 125 | 130 | 142.8 | 146.1 | 143 | 132 | 118.2526 | 146 | 143.2 | 143 | 126.3 | 136 | 140.5569 | 141 |  |  |  | 16
| 1 | 4 | 150.5 | 142 | 141.5 | 143.5 | 138 | 160 | 129 | 139.1 | 154.5 | 137.2 | 130 | 118.2021 | 144 | 143.5 | 139 | 122.8 |  | 147.5125 | 137 |  |  |  | 142.1221
| 1 | 5 | 152 | 139.5 | 139 | 143 | 140.5 | 149 | 133 | 139.1 | 169.8 | 144.8 | 131 | 117.1364 | 143 | 144.5 | 132 | 136.5 | 136 | 139.4355 | 138 |  |  |  | 113
| 1 | 6 | 155 | 149 | 137 | 146 | 145 | 136 | 133 | 146.7 | 156 | 146.9 | 134 | 117.779 | 145 | 146.2 | 141 | 146.7 | 134 | 145.9148 | 138
| Mean |  | 151.0833 | 142.0833 | 143 | 143.25 | 139.5833 | 147.5 | 131.8333 | 142.9333 | 154.5333 | 140.7 | 131.5 | 117.9502 | 143.6667 | 144.9167 | 140.6667 | 130.6167 | 134.2 | 141.5238 | 138.8333
| Median |  | 151 | 141.25 | 142 | 143.25 | 138.75 | 150 | 131.5 | 142.95 | 153.05 | 141.2 | 131 | 118.1578 | 144 | 144.55 | 142 | 126.45 | 135 | 140.8468 | 138
| Std Dev. |  | 2.3962 | 3.5695 | 5.0398 | 2.0187 | 2.9566 | 14.7071 | 2.6394 | 3.4227 | 8.2873 | 5.2023 | 1.3784 | 0.4347 | 2.0656 | 1.6461 | 4.761 | 9.1977 | 2.49 | 4.6631 | 1.7224
| Rel.Std.Dev. |  | 1.586% | 2.512% | 3.524% | 1.409% | 2.118% | 9.971% | 2.002% | 2.395% | 5.363% | 3.697% | 1.048% | 0.3685% | 1.438% | 1.136% | 3.385% | 7.042% | 1.855% | 3.295% | 1.241%
| PDM3 |  | 6.305% | -0.02728% | 0.6177% | 0.7936% | -1.786% | 3.784% | -7.239% | 0.5708% | 8.733% | -1.001% | -7.474% | -17.01% | 1.087% | 1.966% | -1.024% | -8.095% | -5.574% | -0.421% | -2.314%
| Z-Score (Absolute) |  | 2.4 | 0.14 | 0.37 | 0.43 | 0.49 | 1.5 | 2.43 | 0.35 | 3.26 | 0.21 | 2.51 | 5.91 | 0.54 | 0.85 | 0.21 | 2.73 | 1.84 | 0 | 0.67
| Analytical results for Re in OREAS 926b (Indicative Value 1.72 ppb) | [23 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 20 | [9 empty] | ppb
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1 | < 2 | 2 | 0.9 | 3 | < 1 | 6 | < 20 | 2 | < 3 | < 50 | < 2 | < 2 | [9 empty] | 1
| 1 | 2 | 1.1 | < 2 | < 2 | 0.5 | 3 | 1 | 5 | < 20 | 2 | < 3 | < 50 | < 2 | < 2 | [9 empty] | 12
| 1 | 3 | 0.9 | < 2 | < 2 | 0.5 | 3 | 2 | 7 | < 20 | < 2 | 3 | < 50 | < 2 | 2 | [9 empty] | 16
| 1 | 4 | 0.9 | 2 | < 2 | 0.7 | 3 | < 1 | 5 | < 20 | 2 | 3 | < 50 | < 2 | [10 empty] | 1.7185
| 1 | 5 | 1.3 | < 2 | < 2 | 0.6 | 3 | 2 | 5 | < 20 | 3 | 3 | < 50 | < 2 | 5 | [9 empty] | 23
| 1 | 6 | 1.2 | < 2 | < 2 | 0.7 | 4 | 1 | 6 | < 20 | 4 | 3 | < 50 | < 2 | 7
| Mean |  | 1.0667 | 2 | 2 | 0.65 | 3.1667 | 1.5 | 5.6667 |  | 2.6 | 3 |  |  | 4.6667
| Median |  | 1.05 | 2 | 2 | 0.65 | 3 | 1.5 | 5.5 |  | 2 | 3 |  |  | 5
| Std Dev. |  | 0.1633 |  |  | 0.15166 | 0.40825 | 0.57735 | 0.8165 |  | 0.89443 | 0 |  |  | 2.5166
| Rel.Std.Dev. |  | 15.31% |  |  | 23.33% | 12.89% | 38.49% | 14.41% |  | 34.4% | 0% |  |  | 53.93%
| PDM3 |  | -37.93% | 16.38% | 16.38% | -62.18% | 84.27% | -12.72% | 229.7% |  | 51.29% | 74.57% |  |  | 171.6%
| Z-Score (Absolute) |  | 0.73 | 0.67 | 0.67 | 0.97 | 0.48 | 0.67 | 1.93 | 4.43 | 0 | 0.1 | 13.1 | 0.77 | 0.5
| Analytical results for S in OREAS 926b (Certified Value 0.988 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.99381 | 0.97 | 1.02 | 1 | 0.99 | 0.98 | 0.88 | 0.99 | 1.03 | 0.92 | 0.9628 | 0.952 | 1.0834 | 0.985 | 1.02 | 1.02 | 0.92 | 1.008 | 1 | 0.872 | 1.01 |  | 1
| 1 | 2 | 0.98388 | 0.98 | 1.02 | 1.01 | 1 | 0.97 | 1 | 0.99 | 1.02 | 0.9 | 0.9679 | 0.965 | 1.0486 | 0.975 | 1 | 1.03 | 0.89 | 0.994 | 0.97 | 0.884 | 0.989 |  | #N/A
| 1 | 3 | 0.98079 | 0.98 | 1 | 0.97 | 1 | 0.96 | 0.93 | 0.98 | 1.04 | 0.9 | 0.9856 | 0.957 | 1.0478 | 1 | 1 | 1.01 | 0.91 | 0.978 | 1 | 0.88 | 0.986 |  | 16
| 1 | 4 | 0.97062 | 0.96 | 1.04 | 0.99 | 1 | 0.98 | 1.04 | 0.97 | 1.02 | 0.94 | 0.9436 | 0.966 | 1.0521 | 0.99 | 1.01 | 1.05 | 0.93 | 1.016 |  | 0.871 | 0.991 |  | 0.98809
| 1 | 5 | 0.9912 | 1 | 1.01 | 1 | 1 | 0.97 | 0.89 | 0.98 | 0.99 | 0.94 | 0.9946 | 0.951 | 1.0557 | 1 | 1.01 | 1 | 1 | 0.995 | 0.99 | 0.834 | 0.996 |  | 114
| 1 | 6 | 0.99075 | 1 | 1.04 | 0.96 | 0.99 | 0.99 | 0.89 | 0.97 | 1.01 | 1 | 0.9549 | 0.992 | 1.0409 | 1.01 | 1.03 | 1.06 | 1.04 | 1.008 | 0.99 | 0.894 | 0.972
| Mean |  | 0.98518 | 0.98167 | 1.0217 | 0.98833 | 0.99667 | 0.975 | 0.93833 | 0.98 | 1.0183 | 0.93333 | 0.96823 | 0.96383 | 1.0548 | 0.99333 | 1.0117 | 1.0283 | 0.94833 | 0.99983 | 0.99 | 0.8725 | 0.99067
| Median |  | 0.98732 | 0.98 | 1.02 | 0.995 | 1 | 0.975 | 0.91 | 0.98 | 1.02 | 0.93 | 0.96535 | 0.961 | 1.0504 | 0.995 | 1.01 | 1.025 | 0.925 | 1.0015 | 0.99 | 0.876 | 0.99
| Std Dev. |  | 0.0086578 | 0.016021 | 0.016021 | 0.019408 | 0.005164 | 0.010488 | 0.066758 | 0.0089443 | 0.017224 | 0.037238 | 0.019032 | 0.015171 | 0.014869 | 0.012517 | 0.01169 | 0.023166 | 0.058452 | 0.01363 | 0.012247 | 0.020666 | 0.012453
| Rel.Std.Dev. |  | 0.8788% | 1.632% | 1.568% | 1.964% | 0.5181% | 1.076% | 7.115% | 0.9127% | 1.691% | 3.99% | 1.966% | 1.574% | 1.41% | 1.26% | 1.156% | 2.253% | 6.164% | 1.363% | 1.237% | 2.369% | 1.257%
| PDM3 |  | -0.2948% | -0.6499% | 3.398% | 0.02476% | 0.8681% | -1.325% | -5.036% | -0.8186% | 3.061% | -5.542% | -2.009% | -2.455% | 6.747% | 0.5308% | 2.386% | 4.073% | -4.023% | 1.189% | 0.1934% | -11.7% | 0.2609%
| Z-Score (Absolute) |  | 0.11 | 0.22 | 1.12 | 0 | 0.28 | 0.45 | 1.68 | 0.28 | 1.01 | 1.85 | 0.67 | 0.82 | 2.23 | 0.17 | 0.78 | 1.34 | 1.34 | 0.39 | 0.06 | 3.89 | 0.08
| Analytical results for Sb in OREAS 926b (Certified Value 4.63 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 15 | 5.5 | 4.94 | 4.83 | 4.5 | 4.71 | 4.83 | 1.2 | 6.18 | 4.7 | 5.01 | 4.06 | 4.6 | 4.1295 | 4.7 | 4.3 | 2.59 | 4.4 | < 10 | 5.35 | < 10 |  | 1
| 1 | 2 | 14 | 5.07 | 4.83 | 4.76 | 4.61 | 4.57 | 4.37 | 1.3 | 8.82 | 4.4 | 5.09 | 4.19 | 4.6 | 3.9895 | 4.8 | 4.2 | 2.65 | 4.24 | < 10 | 4.87 | < 10 |  | 9
| 1 | 3 | 15 | 5.24 | 4.85 | 4.63 | 4.33 | 4.59 | 4.24 | 2.1 | 3.33 | 4.2 | 4.99 | 4.25 | 4.6 | 4.0474 | 4.9 | 4.4 | 2.57 | 4.46 | < 10 | 5.15 | < 10 |  | 16
| 1 | 4 | 15.5415 | 4.66 | 4.92 | 4.69 | 4.32 | 4.74 | 4.37 | 1.9 | 5.04 | 4.6 | 5.29 | 4.01 | 4.5 | 3.8836 | 4.8 | 4.4 | 2.45 | 4.47 | < 10 |  | < 10 |  | 4.6262
| 1 | 5 | 17 | 5.22 | 4.96 | 4.85 | 4.38 | 4.75 | 4.46 | 2 | 7.75 | 4.6 | 5.49 | 4.38 | 4.4 | 3.9588 | 4.9 | 4.2 | 2.51 | 4.37 | < 10 | 4.89 | < 10 |  | 115
| 1 | 6 | 16 | 6.68 | 4.89 | 4.96 | 4.25 | 4.78 | 4.55 | 3.4 | 6.72 | 4.9 | 5.34 | 4.4 | 4.6 | 3.9324 | 4.9 | 4.3 | 2.48 | 4.49 | < 10 | 4.89 | < 10
| Mean |  | 15.4236 | 5.395 | 4.8983 | 4.7867 | 4.3983 | 4.69 | 4.47 | 1.9833 | 6.3067 | 4.5667 | 5.2017 | 4.215 | 4.55 | 3.9902 | 4.8333 | 4.3 | 2.5417 | 4.405 |  | 5.03
| Median |  | 15.2708 | 5.23 | 4.905 | 4.795 | 4.355 | 4.725 | 4.415 | 1.95 | 6.45 | 4.6 | 5.19 | 4.22 | 4.6 | 3.9742 | 4.85 | 4.3 | 2.54 | 4.43 |  | 4.89
| Std Dev. |  | 1.0214 | 0.68742 | 0.051153 | 0.11877 | 0.13288 | 0.088318 | 0.20445 | 0.78846 | 1.952 | 0.24221 | 0.20203 | 0.16084 | 0.083666 | 0.087584 | 0.08165 | 0.089443 | 0.074944 | 0.092682 |  | 0.21307
| Rel.Std.Dev. |  | 6.623% | 12.74% | 1.044% | 2.481% | 3.021% | 1.883% | 4.574% | 39.75% | 30.95% | 5.304% | 3.884% | 3.816% | 1.839% | 2.195% | 1.689% | 2.08% | 2.949% | 2.104% |  | 4.236%
| PDM3 |  | 233.4% | 16.62% | 5.882% | 3.468% | -4.926% | 1.379% | -3.377% | -57.13% | 36.32% | -1.287% | 12.44% | -8.889% | -1.647% | -13.75% | 4.477% | -7.051% | -45.06% | -4.782% |  | 8.728%
| Z-Score (Absolute) |  | 23.35 | 1.53 | 0.45 | 0.21 | 0.63 | 0 | 0.48 | 5.89 | 3.52 | 0.27 | 1.11 | 1.03 | 0.3 | 1.52 | 0.31 | 0.85 | 4.67 | 0.62 | 0.67 | 0.74 | 0.67
| Analytical results for Sc in OREAS 926b (Certified Value 9.69 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 9.24 | 9.8 | 9.8 | 9.22 | 9.6 | 9 | 7.7 | < 20 | 9.9 | 9.9 | 9 | 8.4036 | 10 | 10 | 9.6 | 8.2 | 10.3 | 8.4867 | 10.3 |  |  |  | 1
| 1 | 2 | 9.51 | 9.7 | 10.8 | 9.04 | 9.3 | 10 | 7.9 | < 20 | 10.8 | 9.89 | 9 | 8.288 | 10 | 10 | 9.8 | 6.7 | 9.6 | 9.0453 | 10 |  |  |  | 13
| 1 | 3 | 9.52 | 9.6 | 11 | 9.22 | 9.1 | 10 | 7.7 | < 20 | 10.6 | 10.09 | 9 | 8.4529 | 10 | 9 | 9.8 | 7 | 9.9 | 8.6413 | 10.1 |  |  |  | 16
| 1 | 4 | 9.76 | 9.7 | 10.4 | 9.31 | 9.5 | 10 | 7.7 | < 20 | 11.8 | 9.75 | 9 | 8.3714 | 10 | 9 | 10 | 7.2 |  | 8.5004 | 10.1 |  |  |  | 9.692
| 1 | 5 | 9.78 | 9.6 | 10.9 | 9.17 | 9.5 | 9 | 7.9 | < 20 | 11.4 | 10.07 | 9 | 8.515 | 10 | 10 | 9.4 | 8.7 | 10.3 | 8.342 | 10.4 |  |  |  | 116
| 1 | 6 | 9.84 | 9.9 | 10.1 | 9.34 | 9.7 | 9 | 7.8 | < 20 | 11.4 | 9.96 | 9 | 8.5641 | 10 | 10 | 9.5 | 9.1 | 9.9 | 8.9002 | 9.8
| Mean |  | 9.6083 | 9.7167 | 10.5 | 9.2167 | 9.45 | 9.5 | 7.7833 |  | 10.9833 | 9.9433 | 9 | 8.4325 | 10 | 9.6667 | 9.6833 | 7.8167 | 10 | 8.6526 | 10.1167
| Median |  | 9.64 | 9.7 | 10.6 | 9.22 | 9.5 | 9.5 | 7.75 |  | 11.1 | 9.93 | 9 | 8.4282 | 10 | 10 | 9.7 | 7.7 | 9.9 | 8.5708 | 10.1
| Std Dev. |  | 0.22772 | 0.1169 | 0.48166 | 0.10708 | 0.21679 | 0.54772 | 0.098319 |  | 0.68823 | 0.12644 | 0 | 0.099997 | 0 | 0.5164 | 0.22286 | 0.98675 | 0.3 | 0.26938 | 0.2137
| Rel.Std.Dev. |  | 2.37% | 1.203% | 4.587% | 1.162% | 2.294% | 5.766% | 1.263% |  | 6.266% | 1.272% | 0% | 1.186% | 0% | 5.342% | 2.301% | 12.62% | 3% | 3.113% | 2.112%
| PDM3 |  | -0.8628% | 0.255% | 8.337% | -4.904% | -2.496% | -1.981% | -19.69% |  | 13.32% | 2.594% | -7.139% | -12.99% | 3.178% | -0.2609% | -0.08897% | -19.35% | 3.178% | -10.72% | 4.382%
| Z-Score (Absolute) |  | 0.12 | 0.05 | 1.27 | 0.73 | 0.36 | NA | 2.96 | 0.49 | 2.02 | 0.4 | NA | 1.95 | NA | NA | 0 | 2.9 | 0.49 | 1.6 | 0.67
Results from laboratories 8, 13, 15 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 926b (Certified Value 11.8 ppm) | [23 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 11.9 | 13 | 12 | 11.65 | 13 | 13.7 | 7 | 11 | 7 | 8.7 | 12 | 10 | 12 | 11.92 | 11 | 10.49 | 13.3 | 3.9 |  |  |  |  | 1
| 1 | 2 | 12.5 | 12 | 11 | 11.65 | 12 | 11.7 | 6 | 15 | 7.8 | 9.21 | 12 | 10 | 12 | 11.92 | 11 | 11.397 | 13 | 4.0906 |  |  |  |  | 14
| 1 | 3 | 12.55 | 13 | 12 | 11.8 | 12 | 13.1 | 8 | 12 | 7.4 | 9.37 | 12 | 10 | 12 | 11.44 | 10 | 11.865 | 13.7 | 4.0185 |  |  |  |  | 16
| 1 | 4 | 12 | 12 | 12 | 11.75 | 12 | 12.2 | 7 | 9 | 8.2 | 8.83 | 12 | 10 | 13 | 12.25 | 11 | 10.605 |  | 3.7414 |  |  |  |  | 11.7677
| 1 | 5 | 12.9 | 12 | 11 | 11.55 | 12 | 12.1 | 7 | 12 | 8.1 | 9.51 | 12 | 10 | 12 | 11.7 | 12 | 11.458 | 12.1 | 4.2848 |  |  |  |  | 117
| 1 | 6 | 12.7 | 12 | 11 | 11.5 | 12 | 12.1 | 6 | 12 | 10 | 9.81 | 11 | 10 | 12 | 11.69 | 12 | 13.827 | 12.1 | 4.2874
| Mean |  | 12.425 | 12.3333 | 11.5 | 11.65 | 12.1667 | 12.4833 | 6.8333 | 11.8333 | 8.0833 | 9.2383 | 11.8333 | 10 | 12.1667 | 11.82 | 11.1667 | 11.607 | 12.84 | 4.0538
| Median |  | 12.525 | 12 | 11.5 | 11.65 | 12 | 12.15 | 7 | 12 | 7.95 | 9.29 | 12 | 10 | 12 | 11.81 | 11 | 11.4275 | 13 | 4.0545
| Std Dev. |  | 0.39465 | 0.5164 | 0.54772 | 0.11402 | 0.40825 | 0.75476 | 0.75277 | 1.9408 | 1.04 | 0.4184 | 0.40825 | 0 | 0.40825 | 0.2759 | 0.75277 | 1.2094 | 0.71972 | 0.21533
| Rel.Std.Dev. |  | 3.176% | 4.187% | 4.763% | 0.9787% | 3.355% | 6.046% | 11.02% | 16.4% | 12.87% | 4.529% | 3.45% | 0% | 3.355% | 2.334% | 6.741% | 10.42% | 5.605% | 5.312%
| PDM3 |  | 5.586% | 4.807% | -2.275% | -1% | 3.39% | 6.081% | -41.93% | 0.5576% | -31.31% | -21.49% | 0.5576% | -15.02% | 3.39% | 0.4443% | -5.108% | -1.366% | 9.112% | -65.55%
| Z-Score (Absolute) |  | 0.8 | 0.69 | 0.27 | 0.1 | 0.5 | 0.87 | 5.67 | 0.11 | 4.22 | 2.89 | 0.11 | 2.01 | 0.5 | 0.1 | 0.66 | 0.15 | 1.28 | 8.88
| Analytical results for Sm in OREAS 926b (Certified Value 5.56 ppm) | [23 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 6.2 | 5.5 | 6 | 5.29 | 5.047 | 5.6 | 5.5 | 5.3 | [14 empty] | 1
| 1 | 2 | 5.5 | 5.6 | 6 | 5.26 | 5.134 | 5.55 | 5.8 | 5.3 | [14 empty] | 1
| 1 | 3 | 5.8 | 5.3 | 6.1 | 5.75 | 5.5423 | 5.65 | 5.75 | 5.3 | [14 empty] | 16
| 1 | 4 | 5.9 | 5.5 | 5.9 | 5.58 | 4.9042 | 5.55 | 5.15 | 5.3 | [14 empty] | 5.5611
| 1 | 5 | 6 | 5.2 | 6.6 | 5.77 | 5.0521 | 5.55 | 5.34 | 5.4 | [14 empty] | 118
| 1 | 6 | 5.8 | 5.4 | 6 | 5.59 | 5.002 | 5.9 | 5.47 | 5.3
| Mean |  | 5.8667 | 5.4167 | 6.1 | 5.54 | 5.1136 | 5.6333 | 5.5017 | 5.3167
| Median |  | 5.85 | 5.45 | 6 | 5.585 | 5.0495 | 5.575 | 5.485 | 5.3
| Std Dev. |  | 0.23381 | 0.1472 | 0.25298 | 0.22 | 0.22305 | 0.13663 | 0.24555 | 0.040825
| Rel.Std.Dev. |  | 3.985% | 2.717% | 4.147% | 3.971% | 4.362% | 2.425% | 4.463% | 0.7679%
| PDM3 |  | 5.495% | -2.597% | 9.691% | -0.379% | -8.047% | 1.299% | -1.068% | -4.395%
| Z-Score (Absolute) |  | 1.47 | 0.44 | 2.47 | 0.08 | 1.73 | 0.48 | 0.08 | 0.87
| Analytical results for Sn in OREAS 926b (Certified Value 10.5 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 7 | 10.35 | 10.8 | 10.7 | 10.4 | 11 | 11 | 6.8 | 11 | 10.7 | 9.41 | 10.6 | 10.5 | 10.6 | 10.5 | 10.6 | 10.968 | 10.5 | 10.1847 | 10.8 |  |  | 1
| 1 | 2 | 5.85 | 10.5 | 10.7 | 10.1 | 10.1 | 10.2 | 10 | 7.2 | 10 | 10.7 | 9.88 | 10.2 | 10.3 | 10.4 | 10.4 | 10.5 | 13.427 | 10.2 | 10.5311 | 10.7 |  |  | 2
| 1 | 3 | 7.84 | 10.65 | 10.6 | 10.1 | 10.45 | 10.2 | 11 | 6.8 | 11 | 10.5 | 10.06 | 10.2 | 10.8 | 10.2 | 10.47 | 10.7 | 11.334 | 10.7 | 10.3552 | 11.4 |  |  | 16
| 1 | 4 | 8.21 | 11 | 11 | 10.1 | 10.5 | 10.4 | 11 | 6.8 | 10 | 11.2 | 9.36 | 10.3 | 10.3 | 10.3 | 9.82 | 10.3 | 9.958 |  | 10.1734 | 13.3 |  |  | 10.4575
| 1 | 5 | 8.83 | 10.85 | 10.8 | 10.1 | 10.6 | 10.5 | 11 | 7.2 | 10 | 11.9 | 10.03 | 10 | 10.4 | 10.5 | 9.96 | 11.2 | 10.803 | 9.8 | 7.9278 | 10.6 |  |  | 119
| 1 | 6 | 4.24 | 13.65 | 11 | 9.9 | 10.75 | 12 | 11 | 7.1 | 11 | 11.3 | 10.36 | 10.2 | 10.5 | 10.6 | 10.34 | 11.7 | 10.258 | 9.7 | 9.6283 | 10.3
| Mean |  | 6.995 | 11.1667 | 10.8167 | 10.1667 | 10.4667 | 10.7167 | 10.8333 | 6.9833 | 10.5 | 11.05 | 9.85 | 10.25 | 10.4667 | 10.4333 | 10.2483 | 10.8333 | 11.1247 | 10.18 | 9.8001 | 11.1833
| Median |  | 7.42 | 10.75 | 10.8 | 10.1 | 10.475 | 10.45 | 11 | 6.95 | 10.5 | 10.95 | 9.955 | 10.2 | 10.45 | 10.45 | 10.37 | 10.65 | 10.8855 | 10.2 | 10.1791 | 10.75
| Std Dev. |  | 1.7005 | 1.2388 | 0.16021 | 0.27325 | 0.21833 | 0.69402 | 0.40825 | 0.20412 | 0.54772 | 0.52058 | 0.39273 | 0.19748 | 0.18619 | 0.1633 | 0.28653 | 0.52026 | 1.2316 | 0.43243 | 0.9659 | 1.098
| Rel.Std.Dev. |  | 24.31% | 11.09% | 1.481% | 2.688% | 2.086% | 6.476% | 3.768% | 2.923% | 5.216% | 4.711% | 3.987% | 1.927% | 1.779% | 1.565% | 2.796% | 4.802% | 11.07% | 4.248% | 9.856% | 9.818%
| PDM3 |  | -33.11% | 6.781% | 3.434% | -2.781% | 0.08741% | 2.478% | 3.594% | -33.22% | 0.4062% | 5.666% | -5.809% | -1.984% | 0.08741% | -0.2313% | -2% | 3.594% | 6.38% | -2.654% | -6.287% | 6.941%
| Z-Score (Absolute) |  | 6.21 | 1.29 | 0.66 | 0.51 | 0.03 | 0.48 | NA | 6.23 | NA | 1.08 | 1.08 | 0.36 | 0.03 | 0.03 | 0.36 | 0.69 | 1.21 | 0.49 | 1.17 | 1.32
Results from laboratories 8 and 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 926b (Certified Value 40.2 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 39.5084 | 40.53 | 39.6 | 41 | 40.5 | 39.6 | 40.7 | 40.3 | 26.9 | 33.9 | 42 | 46 | 39.5 | 42.13 | 46 | 40 | 40 | 39.2 | 40.557 | 39.3 | 30 | 42 | 1
| 1 | 2 | 38.8751 | 39.94 | 39.5 | 39.6 | 39.2 | 38.6 | 40.5 | 41.7 | 28.6 | 35.3 | 41 | 46 | 40.1 | 42.5 | 46 | 39 | 41 | 34.2 | 40.181 | 37.4 | 30 | 41 | 3
| 1 | 3 | 38.9991 | 40.05 | 40.1 | 39.8 | 40.2 | 39.7 | 39.4 | 41.5 | 27.5 | 33.7 | 41 | 47 | 40.3 | 42.82 | 46 | 39 | 40 | 36.8 | 40.078 | 38.9 | 30 | 42 | 16
| 1 | 4 | 39.628 | 39.21 | 40.6 | 40.5 | 39.6 | 40 | 40.9 | 41.1 | 27.6 | 33 | 43 | 45 | 39.2 | 41.13 | 46 | 39 | 40 | 36.1 | 40.747 |  | 30 | 41 | 40.2433
| 1 | 5 | 39.9557 | 40.69 | 41.1 | 40.2 | 39.1 | 39.8 | 40.8 | 41.9 | 27.7 | 32.9 | 41 | 47 | 39.5 | 42.62 | 46 | 40 | 40 | 40.9 | 39.715 | 37.8 | 30 | 42 | 120
| 1 | 6 | 38.8574 | 40.03 | 41 | 41.6 | 38.4 | 40.6 | 42.2 | 41.2 | 27.2 | 35.3 | 43 | 46 | 40.4 | 42.43 | 48 | 40 | 45 | 42.8 | 40.651 | 37.2 | 30 | 41
| Mean |  | 39.304 | 40.075 | 40.3167 | 40.45 | 39.5 | 39.7167 | 40.75 | 41.2833 | 27.5833 | 34.0167 | 41.8333 | 46.1667 | 39.8333 | 42.2717 | 46.3333 | 39.5 | 41 | 38.3333 | 40.3215 | 38.12 | 30 | 41.5
| Median |  | 39.2538 | 40.04 | 40.35 | 40.35 | 39.4 | 39.75 | 40.75 | 41.35 | 27.55 | 33.8 | 41.5 | 46 | 39.8 | 42.465 | 46 | 39.5 | 40 | 38 | 40.369 | 37.8 | 30 | 41.5
| Std Dev. |  | 0.45779 | 0.52045 | 0.69113 | 0.75299 | 0.76942 | 0.65243 | 0.8961 | 0.56716 | 0.57764 | 1.0666 | 0.98319 | 0.75277 | 0.49666 | 0.60377 | 0.8165 | 0.54772 | 2 | 3.216 | 0.39798 | 0.93113 | 0 | 0.54772
| Rel.Std.Dev. |  | 1.165% | 1.299% | 1.714% | 1.862% | 1.948% | 1.643% | 2.199% | 1.374% | 2.094% | 3.136% | 2.35% | 1.631% | 1.247% | 1.428% | 1.762% | 1.387% | 4.878% | 8.39% | 0.987% | 2.443% | 0% | 1.32%
| PDM3 |  | -2.334% | -0.4181% | 0.1824% | 0.5137% | -1.847% | -1.309% | 1.259% | 2.584% | -31.46% | -15.47% | 3.951% | 14.72% | -1.019% | 5.04% | 15.13% | -1.847% | 1.88% | -4.746% | 0.1944% | -5.276% | -25.45% | 3.123%
| Z-Score (Absolute) |  | 0.61 | 0.08 | 0.08 | 0.17 | 0.47 | 0.33 | 0.38 | 0.74 | 8.59 | 4.21 | 1.12 | 4.07 | 0.25 | 1.41 | 4.18 | 0.47 | 0.55 | 1.27 | 0.09 | 1.41 | 6.95 | 0.89
| Analytical results for Ta in OREAS 926b (Indicative Value 0.32 ppm) | [23 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 0.371 | 0.43 | 0.34 | 0.347 | 0.36 | 0.6 | 0.34 | 0.2 | 0.5 | 0.13 | 0.23 | 0.5 | < 0.5 | 0.21 | 0.48 | 0.23 | 0.47061 | 0.78 |  |  |  |  | 1
| 1 | 2 | 0.377 | 0.41 | 0.32 | 0.329 | 0.39 | 0.7 | 0.33 | 0.2 | 0.4 | 0.18 | 0.21 | 0.5 | < 0.5 | 0.22 | 0.48 | 0.21 | 0.47228 | 0.77 |  |  |  |  | 4
| 1 | 3 | 0.391 | 0.41 | 0.34 | 0.329 | 0.41 | 0.7 | 0.24 | 0.2 | 0.4 | 0.15 | 0.26 | 0.5 | < 0.5 | 0.22 | 0.52 | 0.23 | 0.4863 | 0.79 |  |  |  |  | 16
| 1 | 4 | 0.373 | 0.44 | 0.34 | 0.335 | 0.41 | 0.6 | 0.22 | 0.2 | 0.4 | 0.13 | 0.22 | 0.6 | < 0.5 | 0.22 | 0.43 |  | 0.38127 | 0.77 |  |  |  |  | 0.32188
| 1 | 5 | 0.388 | 0.42 | 0.33 | 0.337 | 0.39 | 0.8 | 0.18 | 0.2 | 0.4 | 0.13 | 0.22 | 0.6 | < 0.5 | 0.2 | 0.56 | 0.25 | 0.2996 | 0.78 |  |  |  |  | 22
| 1 | 6 | 0.36 | 0.43 | 0.33 | 0.349 | 0.42 | 1 | 0.19 | 0.2 | 0.4 | 0.14 | 0.21 | 0.6 | < 0.5 | 0.21 | 0.56 | 0.27 | 0.41141 | 0.77
| Mean |  | 0.37667 | 0.42333 | 0.33333 | 0.33767 | 0.39667 | 0.73333 | 0.25 | 0.2 | 0.41667 | 0.14333 | 0.225 | 0.55 |  | 0.21333 | 0.505 | 0.238 | 0.42024 | 0.77667
| Median |  | 0.375 | 0.425 | 0.335 | 0.336 | 0.4 | 0.7 | 0.23 | 0.2 | 0.4 | 0.135 | 0.22 | 0.55 |  | 0.215 | 0.5 | 0.23 | 0.44101 | 0.775
| Std Dev. |  | 0.011466 | 0.012111 | 0.008165 | 0.008641 | 0.021602 | 0.15055 | 0.069282 | 3.0405e-17 | 0.040825 | 0.019664 | 0.018708 | 0.054772 |  | 0.008165 | 0.051284 | 0.022804 | 0.071778 | 0.008165
| Rel.Std.Dev. |  | 3.044% | 2.861% | 2.449% | 2.559% | 5.446% | 20.53% | 27.71% | 1.52e-14% | 9.798% | 13.72% | 8.315% | 9.959% |  | 3.827% | 10.16% | 9.581% | 17.08% | 1.051%
| PDM3 |  | 17.02% | 31.52% | 3.558% | 4.904% | 23.23% | 127.8% | -22.33% | -37.87% | 29.45% | -55.47% | -30.1% | 70.87% |  | -33.72% | 56.89% | -26.06% | 30.56% | 141.3%
| Z-Score (Absolute) |  | 0.32 | 0.68 | 0.02 | 0.02 | 0.48 | NA | 0.67 | NA | NA | 1.5 | 0.86 | NA | 0.67 | 0.95 | 1.32 | 0.76 | 0.66 | 3.43
Results from laboratories 8, 10, 11 and 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 926b (Certified Value 0.56 ppm) | [23 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 17 | 18 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 0.6 | 0.48 | 0.6 | < 0.2 | 0.6 | 0.51 | 0.5 | 0.51 | [14 empty] | 1
| 1 | 2 | 0.6 | 0.53 | 0.6 | < 0.2 | 0.66 | 0.53 | 0.48 | 0.51 | [14 empty] | 21
| 1 | 3 | 0.7 | 0.53 | 0.6 | < 0.2 | 0.62 | 0.56 | 0.5 | 0.52 | [14 empty] | 16
| 1 | 4 | 0.6 | 0.45 | 0.6 | < 0.2 | 0.6 | 0.6 | 0.47 | 0.49 | [14 empty] | 0.55619
| 1 | 5 | 0.6 | 0.49 | 0.6 | < 0.2 | 0.58 | 0.56 | 0.52 | 0.52 | [14 empty] | 121
| 1 | 6 | 0.6 | 0.53 | 0.6 | < 0.2 | 0.6 | 0.59 | 0.57 | 0.51
| Mean |  | 0.61667 | 0.50167 | 0.6 |  | 0.61 | 0.55833 | 0.50667 | 0.51
| Median |  | 0.6 | 0.51 | 0.6 |  | 0.6 | 0.56 | 0.5 | 0.51
| Std Dev. |  | 0.040825 | 0.033714 | 0 |  | 0.027568 | 0.034303 | 0.03559 | 0.010954
| Rel.Std.Dev. |  | 6.62% | 6.72% | 0% |  | 4.519% | 6.144% | 7.024% | 2.148%
| PDM3 |  | 10.87% | -9.803% | 7.877% |  | 9.675% | 0.3853% | -8.904% | -8.305%
| Z-Score (Absolute) |  | 1.13 | 0.45 | 0.9 | 5.95 | 1.04 | 0.33 | 0.38 | 0.33
| Analytical results for Te in OREAS 926b (Certified Value 0.077 ppm) | [23 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 21 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 10.6722 | 0.068 | 0.08 | 0.07 | 0.07 | 0.12 | < 0.1 | 0.45 | < 0.5 | 0.16 | 0.08 | 0.1 | < 0.2 | 0.22 | < 0.05 | < 10 | 0.07 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 19.9725 | 0.065 | 0.09 | 0.08 | 0.069 | 0.05 | < 0.1 | 0.5 | < 0.5 | < 0.05 | 0.08 | 0.1 | < 0.2 | 0.2 | 0.1 | < 10 | 0.08 | < 0.1 |  |  |  |  | 12
| 1 | 3 | 3.4435 | 0.066 | 0.08 | 0.08 | 0.068 | 0.08 | 0.1 | 0.17 | < 0.5 | 0.09 | 0.09 | 0.1 | < 0.2 | 0.23 | 0.07 | < 10 | 0.09 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 13.3471 | 0.069 | 0.08 | 0.07 | 0.066 | 0.06 | < 0.1 | 0.17 | < 0.5 | < 0.05 | 0.11 | 0.1 | < 0.2 | 0.22 | < 0.05 | < 10 |  | < 0.1 |  |  |  |  | 0.077259
| 1 | 5 | 11.584 | 0.075 | 0.1 | 0.06 | 0.066 | < 0.05 | < 0.1 | 0.53 | < 0.5 | < 0.05 | 0.1 | 0.1 | < 0.2 | 0.23 | 0.05 | < 10 | 0.08 | < 0.1 |  |  |  |  | 122
| 1 | 6 | 10.0413 | 0.08 | 0.09 | 0.06 | 0.063 | 0.05 | 0.2 | 0.51 | < 0.5 | 0.15 | 0.06 | 0.1 | < 0.2 | 0.23 | 0.07 | < 10 | 0.08 | < 0.1
| Mean |  | 11.5101 | 0.0705 | 0.086667 | 0.07 | 0.067 | 0.072 | 0.15 | 0.38833 |  | 0.13333 | 0.086667 | 0.1 |  | 0.22167 | 0.0725 |  | 0.08
| Median |  | 11.1281 | 0.0685 | 0.085 | 0.07 | 0.067 | 0.06 | 0.15 | 0.475 |  | 0.15 | 0.085 | 0.1 |  | 0.225 | 0.07 |  | 0.08
| Std Dev. |  | 5.3464 | 0.0058224 | 0.008165 | 0.0089443 | 0.0025298 | 0.029496 | 0.070711 | 0.17116 |  | 0.037859 | 0.017512 | 1.5202e-17 |  | 0.01169 | 0.020616 |  | 0.0070711
| Rel.Std.Dev. |  | 46.45% | 8.259% | 9.421% | 12.78% | 3.776% | 40.97% | 47.14% | 44.08% |  | 28.39% | 20.21% | 1.52e-14% |  | 5.274% | 28.44% |  | 8.839%
| PDM3 |  | 14798% | -8.749% | 12.18% | -9.396% | -13.28% | -6.807% | 94.15% | 402.6% |  | 72.58% | 12.18% | 29.43% |  | 186.9% | -6.16% |  | 3.547%
| Z-Score (Absolute) |  | 462.31 | 0.52 | 0.13 | 0.54 | 0.66 | 0.78 | 0 | 12.34 | 6.74 | 0.17 | 0.13 | NA | 0.67 | 5.6 | 1.08 | 198.92 | 0.13 | 1.35
Results from laboratory 13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 926b (Certified Value 12.4 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 11.65 | 13.3 | 12.5 | 12 | 11.65 | 11.6 | 13.7 | 12.2 | 12.2 | 10.33 | 13.1 | 12.3398 | 13.4 | 13.1 | 11.67 | 10.7 | 12.7 | 12.7741 | 12.9 |  |  |  | 1
| 1 | 2 | 11.65 | 13.35 | 13 | 12 | 11.85 | 13.1 | 14.3 | 12.2 | 12.3 | 10.22 | 13.2 | 12.3788 | 13.1 | 12.4 | 11.84 | 10.3 | 12.3 | 12.751 | 13.1 |  |  |  | 5
| 1 | 3 | 11.95 | 13.25 | 13.75 | 12.25 | 11.9 | 11.3 | 13.3 | 12.1 | 12.1 | 11.07 | 13 | 12.3375 | 13.6 | 12.5 | 11.62 | 10.7 | 13 | 11.6644 | 13.2 |  |  |  | 16
| 1 | 4 | 11.85 | 13.65 | 13.6 | 12.3 | 11.75 | 13.1 | 13 | 11.8 | 12.1 | 10.72 | 13.4 | 12.1622 | 13.4 | 12.6 | 11.66 | 10.5 |  | 11.7748 | 12.7 |  |  |  | 12.3984
| 1 | 5 | 11.95 | 13.35 | 13.7 | 12.45 | 11.9 | 12.7 | 13.4 | 11.9 | 13 | 11.03 | 13 | 12.8055 | 13.2 | 12.9 | 11.59 | 11.8 | 12.5 | 11.0092 | 13 |  |  |  | 123
| 1 | 6 | 11.95 | 13.6 | 13.75 | 12.7 | 11.9 | 13.1 | 13.4 | 12.2 | 12.2 | 10.69 | 13.4 | 12.673 | 13.7 | 12.7 | 11.55 | 12.5 | 12.2 | 11.8036 | 12.8
| Mean |  | 11.8333 | 13.4167 | 13.3833 | 12.2833 | 11.825 | 12.4833 | 13.5167 | 12.0667 | 12.3167 | 10.6767 | 13.1833 | 12.4495 | 13.4 | 12.7 | 11.655 | 11.0833 | 12.54 | 11.9629 | 12.95
| Median |  | 11.9 | 13.35 | 13.65 | 12.275 | 11.875 | 12.9 | 13.4 | 12.15 | 12.2 | 10.705 | 13.15 | 12.3593 | 13.4 | 12.65 | 11.64 | 10.7 | 12.5 | 11.7892 | 12.95
| Std Dev. |  | 0.1472 | 0.16633 | 0.51833 | 0.26957 | 0.10368 | 0.82077 | 0.4446 | 0.17512 | 0.34303 | 0.34938 | 0.18348 | 0.24034 | 0.22804 | 0.26077 | 0.10095 | 0.86814 | 0.32094 | 0.6839 | 0.18708
| Rel.Std.Dev. |  | 1.244% | 1.24% | 3.873% | 2.195% | 0.8768% | 6.575% | 3.289% | 1.451% | 2.785% | 3.272% | 1.392% | 1.931% | 1.702% | 2.053% | 0.8661% | 7.833% | 2.559% | 5.717% | 1.445%
| PDM3 |  | -4.557% | 8.213% | 7.944% | -0.9278% | -4.625% | 0.6853% | 9.02% | -2.675% | -0.659% | -13.89% | 6.331% | 0.4121% | 8.079% | 2.433% | -5.996% | -10.61% | 1.142% | -3.513% | 4.449%
| Z-Score (Absolute) |  | 0.67 | 1.06 | 1.02 | 0.18 | 0.68 | 0.04 | 1.17 | 0.42 | 0.15 | 1.94 | 0.8 | 0 | 1.04 | 0.27 | 0.87 | 1.5 | 0.1 | 0.53 | 0.55
| Analytical results for Ti in OREAS 926b (Certified Value 0.203 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | wt.%
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.2371 | 0.24684 | 0.194 | 0.211 | 0.198 | 0.2 | 0.204 | 0.252 | 0.185 | 0.1644 | 0.202 | 0.1351 | 0.193 | 0.20882 | 0.23 | 0.194 | 0.17 | 0.23 | 0.22385 | 0.19 | 0.167 | 0.276 | 1
| 1 | 2 | 0.22539 | 0.23945 | 0.196 | 0.208 | 0.199 | 0.194 | 0.212 | 0.239 | 0.195 | 0.1568 | 0.19 | 0.1349 | 0.194 | 0.20221 | 0.235 | 0.187 | 0.17 | 0.23 | 0.22248 | 0.18 | 0.168 | 0.27 | 22
| 1 | 3 | 0.2364 | 0.2454 | 0.199 | 0.205 | 0.197 | 0.195 | 0.212 | 0.3 | 0.197 | 0.1546 | 0.189 | 0.1382 | 0.188 | 0.20578 | 0.245 | 0.19 | 0.18 | 0.24 | 0.22282 | 0.19 | 0.162 | 0.276 | 16
| 1 | 4 | 0.23628 | 0.24036 | 0.206 | 0.216 | 0.196 | 0.197 | 0.214 | 0.235 | 0.215 | 0.1613 | 0.195 | 0.1333 | 0.186 | 0.20756 | 0.245 | 0.19 | 0.18 | 0.23 | 0.22059 |  | 0.162 | 0.27 | 0.20256
| 1 | 5 | 0.23657 | 0.24673 | 0.207 | 0.208 | 0.199 | 0.198 | 0.213 | 0.254 | 0.2 | 0.152 | 0.188 | 0.136 | 0.192 | 0.20606 | 0.25 | 0.192 | 0.17 | 0.23 | 0.22674 | 0.19 | 0.162 | 0.276 | 124
| 1 | 6 | 0.2255 | 0.24477 | 0.2 | 0.219 | 0.195 | 0.2 | 0.218 | 0.263 | 0.202 | 0.159 | 0.21 | 0.1357 | 0.203 | 0.20963 | 0.26 | 0.195 | 0.19 | 0.25 | 0.22166 | 0.19 | 0.165 | 0.276
| Mean |  | 0.23287 | 0.24393 | 0.20033 | 0.21117 | 0.19733 | 0.19733 | 0.21217 | 0.25717 | 0.199 | 0.15802 | 0.19567 | 0.13553 | 0.19267 | 0.20668 | 0.24417 | 0.19133 | 0.17667 | 0.235 | 0.22302 | 0.188 | 0.16433 | 0.274
| Median |  | 0.23634 | 0.24509 | 0.1995 | 0.2095 | 0.1975 | 0.1975 | 0.2125 | 0.253 | 0.1985 | 0.1579 | 0.1925 | 0.1354 | 0.1925 | 0.20681 | 0.245 | 0.191 | 0.175 | 0.23 | 0.22265 | 0.19 | 0.1635 | 0.276
| Std Dev. |  | 0.0057613 | 0.0032251 | 0.0052409 | 0.0053448 | 0.001633 | 0.0025033 | 0.0045789 | 0.023353 | 0.0098184 | 0.0045133 | 0.0087331 | 0.0016083 | 0.0059217 | 0.002653 | 0.010685 | 0.0029439 | 0.008165 | 0.0083666 | 0.0021274 | 0.0044721 | 0.0027325 | 0.0030984
| Rel.Std.Dev. |  | 2.474% | 1.322% | 2.616% | 2.531% | 0.8275% | 1.269% | 2.158% | 9.081% | 4.934% | 2.856% | 4.463% | 1.187% | 3.074% | 1.284% | 4.376% | 1.539% | 4.622% | 3.56% | 0.9539% | 2.379% | 1.663% | 1.131%
| PDM3 |  | 14.96% | 20.42% | -1.101% | 4.248% | -2.582% | -2.582% | 4.741% | 26.96% | -1.759% | -21.99% | -3.404% | -33.09% | -4.885% | 2.031% | 20.54% | -5.544% | -12.78% | 16.01% | 10.1% | -7.189% | -18.87% | 35.27%
| Z-Score (Absolute) |  | 1.26 | 1.68 | 0.03 | 0.44 | 0.09 | 0.09 | 0.47 | 2.18 | 0.03 | 1.58 | 0.15 | 2.44 | 0.27 | 0.27 | 1.69 | 0.32 | 0.87 | 1.34 | 0.89 | 0.44 | 1.34 | 2.81
| Analytical results for Tl in OREAS 926b (Certified Value 0.94 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 19.2265 | 0.886 | 0.93 | 0.96 | 0.913 | 0.95 | 0.99 | 1.05 | 1.1 | 0.99 | 0.818 | 0.95 | 0.74057 | 0.9 | 0.8 | 1 | < 10 | 1.02 | 0.96806 |  |  |  | 1
| 1 | 2 | 16.1768 | 0.895 | 0.93 | 1 | 0.886 | 0.9 | 0.95 | 1.06 | 1 | 1.03 | 0.791 | 0.96 | 0.66929 | 0.8 | 0.82 | 0.97 | < 10 | 0.92 | 0.95754 |  |  |  | #N/A
| 1 | 3 | 16 | 0.909 | 0.92 | 0.93 | 0.892 | 0.91 | 0.92 | 1.03 | 1.1 | 0.97 | 0.863 | 0.97 | 0.64794 | 0.8 | 0.86 | 1 | < 10 | 0.96 | 0.94974 |  |  |  | 16
| 1 | 4 | 16.326 | 0.888 | 0.94 | 0.93 | 0.891 | 0.89 | 0.96 | 0.98 | 1 | 1.07 | 0.836 | 0.99 | 0.56118 | 0.7 | 0.82 | 0.98 | < 10 |  | 0.98461 |  |  |  | 0.9405
| 1 | 5 | 18 | 0.91 | 0.95 | 0.93 | 0.894 | 0.9 | 0.94 | 1 | 1 | 1.11 | 0.88 | 0.95 | 0.62678 | 0.8 | 0.78 | 0.98 | < 10 | 1.01 | 0.79549 |  |  |  | 125
| 1 | 6 | 17 | 0.911 | 0.94 | 0.93 | 0.912 | 0.96 | 0.96 | 1 | 1 | 1.04 | 0.836 | 0.98 | 0.61912 | 0.7 | 0.83 | 1.03 | < 10 | 1.04 | 0.91587
| Mean |  | 17.1215 | 0.89983 | 0.935 | 0.94667 | 0.898 | 0.91833 | 0.95333 | 1.02 | 1.0333 | 1.035 | 0.83733 | 0.96667 | 0.64414 | 0.78333 | 0.81833 | 0.99333 |  | 0.99 | 0.92855
| Median |  | 16.663 | 0.902 | 0.935 | 0.93 | 0.893 | 0.905 | 0.955 | 1.015 | 1 | 1.035 | 0.836 | 0.965 | 0.63736 | 0.8 | 0.82 | 0.99 |  | 1.01 | 0.95364
| Std Dev. |  | 1.265 | 0.011548 | 0.010488 | 0.028752 | 0.011541 | 0.029269 | 0.023381 | 0.031623 | 0.05164 | 0.051284 | 0.031633 | 0.01633 | 0.059589 | 0.075277 | 0.027142 | 0.021602 |  | 0.04899 | 0.069077
| Rel.Std.Dev. |  | 7.388% | 1.283% | 1.122% | 3.037% | 1.285% | 3.187% | 2.453% | 3.1% | 4.997% | 4.955% | 3.778% | 1.689% | 9.251% | 9.61% | 3.317% | 2.175% |  | 4.948% | 7.439%
| PDM3 |  | 1720% | -4.324% | -0.5848% | 0.6557% | -4.519% | -2.357% | 1.365% | 8.453% | 9.871% | 10.05% | -10.97% | 2.782% | -31.51% | -16.71% | -12.99% | 5.618% |  | 5.263% | -1.27%
| Z-Score (Absolute) |  | 232.89 | 0.67 | 0.17 | 0 | 0.7 | 0.41 | 0.1 | 1.06 | NA | 1.27 | 1.57 | 0.29 | 4.36 | NA | 1.85 | 0.67 | 58.36 | 0.62 | 0.26
Results from laboratories 10 and 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 926b (Indicative Value 0.2 ppm) | [23 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 17 | [16 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [16 empty] | 2
2
| 1 | 1 | 0.3 | 0.2 | 0.2 | 0.14 | 0.2 | 0.16 | [16 empty] | 1
| 1 | 2 | 0.3 | 0.2 | 0.2 | 0.14 | 0.2 | 0.19 | [16 empty] | 17
| 1 | 3 | 0.3 | 0.2 | 0.2 | 0.15 | 0.2 | 0.19 | [16 empty] | 16
| 1 | 4 | 0.3 | 0.2 | 0.2 | 0.17 | 0.2 | 0.19 | [16 empty] | 0.1975
| 1 | 5 | 0.3 | 0.2 | 0.2 | 0.16 | 0.2 | 0.17 | [16 empty] | 23
| 1 | 6 | 0.3 | 0.2 | 0.2 | 0.15 | 0.2 | 0.19
| Mean |  | 0.3 | 0.2 | 0.2 | 0.15167 | 0.2 | 0.18167
| Median |  | 0.3 | 0.2 | 0.2 | 0.15 | 0.2 | 0.19
| Std Dev. |  | 0 | 3.0405e-17 | 3.0405e-17 | 0.01169 | 3.0405e-17 | 0.013292
| Rel.Std.Dev. |  | 0% | 1.52e-14% | 1.52e-14% | 7.708% | 1.52e-14% | 7.316%
| PDM3 |  | 51.9% | 1.266% | 1.266% | -23.21% | 1.266% | -8.017%
| Z-Score (Absolute) |  | 7.36 | 0 | 0 | 3.56 | 0 | 1.35
| Analytical results for U in OREAS 926b (Certified Value 3.97 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 3.68 | 4.1 | 3.9 | 3.77 | 3.7 | 4.2 | 4.42 | 4.1 | 4.2 | 3.43 | 4.11 | 2.8935 | 4 | 4.2 | 3.7 | 3.7 | 3.89 | 4.4125 | 4 |  |  |  | 1
| 1 | 2 | 3.71 | 4.1 | 4.2 | 3.72 | 3.8 | 4 | 4.53 | 4.2 | 4.1 | 3.42 | 4.19 | 2.8852 | 4.1 | 4 | 3.9 | 3.5 | 3.67 | 4.2635 | 3.86 |  |  |  | #N/A
| 1 | 3 | 3.79 | 4 | 3.9 | 3.74 | 3.7 | 4.2 | 4.28 | 4.2 | 4 | 3.69 | 4.1 | 3.0429 | 4.1 | 4 | 3.9 | 3.7 | 3.87 | 3.9891 | 4.1 |  |  |  | 16
| 1 | 4 | 3.74 | 4.1 | 3.8 | 3.7 | 3.7 | 4.1 | 4.29 | 4.3 | 4.2 | 3.58 | 4.06 | 2.9516 | 4 | 4 | 3.9 | 3.6 |  | 4.0731 | 3.94 |  |  |  | 3.9733
| 1 | 5 | 3.77 | 4.2 | 3.9 | 3.77 | 3.8 | 4 | 4.34 | 4.2 | 4.5 | 3.67 | 4.1 | 3.0912 | 4 | 4.1 | 3.8 | 3.8 | 3.97 | 3.7953 | 4.03 |  |  |  | 126
| 1 | 6 | 3.88 | 4.2 | 3.8 | 3.86 | 3.8 | 4 | 4.33 | 4.2 | 4.2 | 3.53 | 4.16 | 2.9903 | 4.2 | 4.1 | 4 | 4.1 | 4.01 | 3.9061 | 4.1
| Mean |  | 3.7617 | 4.1167 | 3.9167 | 3.76 | 3.75 | 4.0833 | 4.365 | 4.2 | 4.2 | 3.5533 | 4.12 | 2.9758 | 4.0667 | 4.0667 | 3.8667 | 3.7333 | 3.882 | 4.0733 | 4.005
| Median |  | 3.755 | 4.1 | 3.9 | 3.755 | 3.75 | 4.05 | 4.335 | 4.2 | 4.2 | 3.555 | 4.105 | 2.971 | 4.05 | 4.05 | 3.9 | 3.7 | 3.89 | 4.0311 | 4.015
| Std Dev. |  | 0.070261 | 0.075277 | 0.1472 | 0.056214 | 0.054772 | 0.098319 | 0.094816 | 0.063246 | 0.16733 | 0.11535 | 0.046904 | 0.081972 | 0.08165 | 0.08165 | 0.10328 | 0.20656 | 0.13161 | 0.2296 | 0.093755
| Rel.Std.Dev. |  | 1.868% | 1.829% | 3.758% | 1.495% | 1.461% | 2.408% | 2.172% | 1.506% | 3.984% | 3.246% | 1.138% | 2.755% | 2.008% | 2.008% | 2.671% | 5.533% | 3.39% | 5.637% | 2.341%
| PDM3 |  | -5.328% | 3.607% | -1.427% | -5.369% | -5.621% | 2.768% | 9.857% | 5.704% | 5.704% | -10.57% | 3.691% | -25.11% | 2.349% | 2.349% | -2.685% | -6.041% | -2.299% | 2.515% | 0.7966%
| Z-Score (Absolute) |  | 1.19 | 0.54 | 0.43 | 1.19 | 1.24 | 0.38 | 1.75 | 0.95 | 0.95 | 2.2 | 0.56 | 5.02 | 0.3 | 0.3 | 0.67 | 1.32 | 0.6 | 0.33 | 0
| Analytical results for V in OREAS 926b (Certified Value 169 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 172.69 | 173.36 | 164.5 | 170 | 170 | 171.5 | 165 | 181 | 146 | 174 | 168 | 188 | 163 | 165.1675 | 160 | 173 | 171 | 181 | 170.462 | 179 | 162.4146 | 180 | 1
| 1 | 2 | 162.69 | 169.48 | 164.5 | 164 | 173 | 168.5 | 170 | 172 | 152 | 178 | 164 | 186 | 166 | 168.2458 | 155 | 170 | 173 | 165 | 168.01 | 157 | 162.7267 | 177 | 6
| 1 | 3 | 174.11 | 172.07 | 164 | 163 | 168 | 168 | 166 | 178 | 143 | 178 | 163 | 191 | 167 | 169.3432 | 160 | 168 | 174 | 184 | 167.015 | 177 | 159.6273 | 179 | 16
| 1 | 4 | 175.85 | 169.1 | 165.5 | 167 | 171 | 169.5 | 167 | 171 | 143 | 175 | 169 | 184 | 164 | 164.6095 | 155 | 169 | 173 | 178 | 171.503 |  | 163.7471 | 177 | 168.955
| 1 | 5 | 176.69 | 174 | 171.5 | 166 | 171 | 171 | 169 | 169 | 145 | 170 | 169 | 190 | 169 | 169.9663 | 155 | 170 | 166 | 187 | 165.674 | 163 | 153.5206 | 180 | 127
| 1 | 6 | 161.74 | 170.4 | 166.5 | 169 | 167 | 172 | 173 | 173 | 142 | 171 | 177 | 188 | 165 | 161.5591 | 165 | 173 | 180 | 193 | 169.722 | 158 | 164.4962 | 178
| Mean |  | 170.6283 | 171.4017 | 166.0833 | 166.5 | 170 | 170.0833 | 168.3333 | 174 | 145.1667 | 174.3333 | 168.3333 | 187.8333 | 165.6667 | 166.4819 | 158.3333 | 170.5 | 172.8333 | 181.3333 | 168.731 | 166.8 | 161.0888 | 178.5
| Median |  | 173.4 | 171.235 | 165 | 166.5 | 170.5 | 170.25 | 168 | 172.5 | 144 | 174.5 | 168.5 | 188 | 165.5 | 166.7066 | 157.5 | 170 | 173 | 182.5 | 168.866 | 163 | 162.5707 | 178.5
| Std Dev. |  | 6.6694 | 2.0507 | 2.8003 | 2.7386 | 2.1909 | 1.6558 | 2.9439 | 4.5607 | 3.656 | 3.3862 | 4.9666 | 2.5626 | 2.1602 | 3.2529 | 4.0825 | 2.0736 | 4.535 | 9.5219 | 2.2123 | 10.4976 | 4.0628 | 1.3784
| Rel.Std.Dev. |  | 3.909% | 1.196% | 1.686% | 1.645% | 1.289% | 0.9735% | 1.749% | 2.621% | 2.519% | 1.942% | 2.95% | 1.364% | 1.304% | 1.954% | 2.578% | 1.216% | 2.624% | 5.251% | 1.311% | 6.294% | 2.522% | 0.7722%
| PDM3 |  | 0.9904% | 1.448% | -1.7% | -1.453% | 0.6185% | 0.6678% | -0.368% | 2.986% | -14.08% | 3.183% | -0.368% | 11.17% | -1.946% | -1.464% | -6.287% | 0.9144% | 2.295% | 7.326% | -0.1326% | -1.275% | -4.656% | 5.649%
| Z-Score (Absolute) |  | 0.28 | 0.45 | 0.72 | 0.63 | 0.14 | 0.16 | 0.23 | 1.01 | 5.29 | 1.09 | 0.23 | 4.04 | 0.81 | 0.63 | 2.41 | 0.25 | 0.76 | 2.62 | 0.14 | 0.56 | 1.81 | 2
| Analytical results for W in OREAS 926b (Certified Value 6.01 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | < 5 | 5.88 | 5.2 | 5.9 | 5.5 | 6.3 | 6 | 6.6 | 6 | 6.1 | 5.26 | 5.7 | 6.5 | 5.8 | < 10 | 5.6 | < 10 | 5.7 | 6.5434 | 7 |  |  | 1
| 1 | 2 | < 5 | 6.04 | 5.1 | 5.6 | 5.59 | 5.9 | 6.3 | 6.79 | 6 | 6.3 | 5.2 | 5.9 | 6.5 | 5.7 | < 10 | 5.4 | < 10 | 5.4 | 6.8969 | 7.1 |  |  | 7
| 1 | 3 | < 5 | 6.08 | 5 | 5.4 | 5.43 | 5.8 | 6.5 | 6.23 | 6 | 6 | 5.68 | 5.9 | 7 | 5.6 | < 10 | 5.6 | < 10 | 5.7 | 6.5312 | 7.3 |  |  | 16
| 1 | 4 | < 5 | 5.95 | 5.4 | 5.4 | 5.53 | 5.8 | 6.2 | 6.17 | 6 | 6.5 | 5.5 | 5.9 | 6.5 | 5.6 | < 10 | 5.5 | < 10 |  | 6.9783 | 7 |  |  | 6.0067
| 1 | 5 | < 5 | 6.14 | 5.1 | 5.5 | 5.53 | 5.9 | 6.5 | 6.08 | 6 | 6.8 | 5.62 | 5.7 | 6.5 | 6 | < 10 | 5.7 | < 10 | 6.3 | 6.3785 | 7.1 |  |  | 128
| 1 | 6 | < 5 | 5.9 | 5.2 | 5.4 | 5.53 | 6 | 6.7 | 6.08 | 6 | 6.5 | 5.55 | 5.8 | 7 | 5.7 | < 10 | 6.1 | < 10 | 6 | 6.4358 | 7.1
| Mean |  |  | 5.9983 | 5.1667 | 5.5333 | 5.5183 | 5.95 | 6.3667 | 6.325 | 6 | 6.3667 | 5.4683 | 5.8167 | 6.6667 | 5.7333 |  | 5.65 |  | 5.82 | 6.6274 | 7.1
| Median |  |  | 5.995 | 5.15 | 5.45 | 5.53 | 5.9 | 6.4 | 6.2 | 6 | 6.4 | 5.525 | 5.85 | 6.5 | 5.7 |  | 5.6 |  | 5.7 | 6.5373 | 7.1
| Std Dev. |  |  | 0.10439 | 0.13663 | 0.19664 | 0.052313 | 0.18708 | 0.25033 | 0.29831 | 0 | 0.29439 | 0.19539 | 0.098319 | 0.2582 | 0.15055 |  | 0.2429 |  | 0.34205 | 0.2493 | 0.10954
| Rel.Std.Dev. |  |  | 1.74% | 2.644% | 3.554% | 0.948% | 3.144% | 3.932% | 4.716% | 0% | 4.624% | 3.573% | 1.69% | 3.873% | 2.626% |  | 4.299% |  | 5.877% | 3.762% | 1.543%
| PDM3 |  |  | -0.1395% | -13.99% | -7.881% | -8.131% | -0.9441% | 5.993% | 5.299% | -0.1117% | 5.993% | -8.963% | -3.164% | 10.99% | -4.551% |  | -5.939% |  | -3.108% | 10.33% | 18.2%
| Z-Score (Absolute) |  | 4.4 | 0.24 | 0.86 | 0.38 | 0.4 | 0.18 | 0.73 | 0.67 | NA | 0.73 | 0.46 | 0 | 1.13 | 0.11 | 1.08 | 0.22 | 1.08 | 0 | 1.08 | 1.7
Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 926b (Certified Value 12.8 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 20 | 22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 1
1
| 1 | 1 | 12.3 | 11.6 | 14.6 | 12.5 | 17 | 19.6 | 9.85 | 11.9 | 13.2 | 11.69 | 14.7 | 10.6586 | 13.6 | 13 | 13.1 | 12 | 17.8 | [5 empty] | 1
| 1 | 2 | 12.55 | 13.7 | 13.5 | 12.7 | 13.7 | 17.3 | 10.8 | 10.9 | 13.2 | 11.7 | 14.6 | 10.7063 | 13.8 | 13 | 12.1 | 11.4 | 17.4 | [5 empty] | #N/A
| 1 | 3 | 12.7 | 15.5 | 14.9 | 12.9 | 13.7 | 20.3 | 10.1 | 10.9 | 13.1 | 12.06 | 15.1 | 10.8266 | 13.7 | 15 | 12.7 | 11.7 | 17.8 | [5 empty] | 16
| 1 | 4 | 13.95 | 13.5 | 14.8 | 13.4 | 14.5 | 17 | 10.6 | 10.9 | 13.2 | 11.58 | 14.6 | 10.5431 | 13.3 | 13 | 11.8 |  | 17.3 | [5 empty] | 12.8213
| 1 | 5 | 13.1 | 13.7 | 14.6 | 12.45 | 13.6 | 17.7 | 10.6 | 10.5 | 14.2 | 12.03 | 15.1 | 10.3774 | 13.8 | 14 | 12.8 | 12.9 | 17.8 | [5 empty] | 129
| 1 | 6 | 13.15 | 14 | 14.3 | 12.8 | 14.5 | 17 | 11.1 | 11 | 13.2 | 11.78 | 14.9 | 10.5391 | 14.1 | 14 | 14.1 | 12.9 | 17.3
| Mean |  | 12.9583 | 13.6667 | 14.45 | 12.7917 | 14.5 | 18.15 | 10.5083 | 11.0167 | 13.35 | 11.8067 | 14.8333 | 10.6085 | 13.7167 | 13.6667 | 12.7667 | 12.18 | 17.5667
| Median |  | 12.9 | 13.7 | 14.6 | 12.75 | 14.1 | 17.5 | 10.6 | 10.9 | 13.2 | 11.74 | 14.8 | 10.6009 | 13.75 | 13.5 | 12.75 | 12 | 17.6
| Std Dev. |  | 0.58431 | 1.2469 | 0.50892 | 0.34412 | 1.2915 | 1.4349 | 0.45871 | 0.46655 | 0.41833 | 0.19552 | 0.23381 | 0.1563 | 0.26394 | 0.8165 | 0.80911 | 0.69065 | 0.2582
| Rel.Std.Dev. |  | 4.509% | 9.123% | 3.522% | 2.69% | 8.907% | 7.906% | 4.365% | 4.235% | 3.134% | 1.656% | 1.576% | 1.473% | 1.924% | 5.974% | 6.338% | 5.67% | 1.47%
| PDM3 |  | 1.068% | 6.593% | 12.7% | -0.2315% | 13.09% | 41.56% | -18.04% | -14.08% | 4.123% | -7.914% | 15.69% | -17.26% | 6.983% | 6.593% | -0.4265% | -5.002% | 37.01%
| Z-Score (Absolute) |  | 0.23 | 0.19 | 0.64 | 0.33 | 0.67 | 2.81 | 1.67 | 1.37 | 0 | 0.9 | 0.87 | 1.61 | 0.21 | 0.19 | 0.34 | 0.69 | 2.47
| Analytical results for Yb in OREAS 926b (Certified Value 1.5 ppm) | [23 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 10 | 11 | 12 | 15 | 17 | 18 | 22 | [14 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [14 empty] | 2
2
| 1 | 1 | 1.9 | 1.2 | 1.5 | 1.11 | 1.65 | 1.38 | 1.4 | 1.68 | [14 empty] | 1
| 1 | 2 | 2 | 1.2 | 1.4 | 1.1 | 1.85 | 1.49 | 1.3 | 1.6 | [14 empty] | 24
| 1 | 3 | 2.1 | 1.1 | 1.5 | 1.18 | 1.8 | 1.39 | 1.4 | 1.67 | [14 empty] | 16
| 1 | 4 | 2 | 1.3 | 1.5 | 1.16 | 1.55 | 1.45 | 1.3 | 1.96 | [14 empty] | 1.5038
| 1 | 5 | 2 | 1.3 | 1.6 | 1.2 | 1.65 | 1.41 | 1.4 | 1.63 | [14 empty] | 130
| 1 | 6 | 2 | 1.2 | 1.5 | 1.15 | 1.6 | 1.45 | 1.6 | 1.68
| Mean |  | 2 | 1.2167 | 1.5 | 1.15 | 1.6833 | 1.4283 | 1.4 | 1.7033
| Median |  | 2 | 1.2 | 1.5 | 1.155 | 1.65 | 1.43 | 1.4 | 1.675
| Std Dev. |  | 0.063246 | 0.075277 | 0.063246 | 0.038987 | 0.1169 | 0.042151 | 0.10954 | 0.12972
| Rel.Std.Dev. |  | 3.162% | 6.187% | 4.216% | 3.39% | 6.945% | 2.951% | 7.825% | 7.616%
| PDM3 |  | 33% | -19.09% | -0.2521% | -23.53% | 11.94% | -5.018% | -6.902% | 13.27%
| Z-Score (Absolute) |  | 1.58 | 0.73 | 0.11 | 0.92 | 0.64 | 0.11 | 0.19 | 0.7
| Analytical results for Zn in OREAS 926b (Certified Value 104 ppm) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 107 | 107.6067 | 156 | 93 | 99 | 116.5 | 100 | 103 | 261 | 122 | 99.8 | 174 | 105 | 105.505 | 110 | 99 | 113 | 97 | 95.78 | 133 | 112.7897 | 120 | 1
| 1 | 2 | 105 | 97.0225 | 103 | 92 | 99 | 100.5 | 105 | 99 | 147 | 104 | 97 | 115 | 102 | 100.983 | 106 | 102 | 107 | 86 | 92.93 | 104 | 116.3749 | 161 | #N/A
| 1 | 3 | 106 | 96.0787 | 102 | 90 | 149 | 99.2 | 100 | 96 | 193 | 98 | 98.1 | 116 | 100 | 101.155 | 102 | 98 | 105 | 88 | 94.02 | 105 | 102.4965 | 113 | 16
| 1 | 4 | 106 | 95.8539 | 100.5 | 92 | 99 | 99.4 | 94 | 104 | 116 | 98 | 98 | 109 | 98 | 104.998 | 100 | 98 | 113 | 84 | 100.2 |  | 103.2266 | 109 | 104.1496
| 1 | 5 | 105 | 100.2697 | 103.5 | 89 | 98 | 98.5 | 95 | 100 | 121 | 100 | 96.7 | 112 | 99 | 100.672 | 112 | 96 | 117 | 100 | 98.78 | 102 | 99.0823 | 145 | 131
| 1 | 6 | 107 | 97.0787 | 101 | 93 | 95 | 100.5 | 96 | 99 | 145 | 99 | 104.6 | 112 | 100 | 102.058 | 104 | 98 | 102 | 91 | 102.3 | 102 | 102.5927 | 109
| Mean |  | 106 | 98.985 | 111 | 91.5 | 106.5 | 102.4333 | 98.3333 | 100.1667 | 163.8333 | 103.5 | 99.0333 | 123 | 100.6667 | 102.5618 | 105.6667 | 98.5 | 109.5 | 91 | 97.335 | 109.2 | 106.0938 | 126.1667
| Median |  | 106 | 97.0506 | 102.5 | 92 | 99 | 99.95 | 98 | 99.5 | 146 | 99.5 | 98.05 | 113.5 | 100 | 101.6065 | 105 | 98 | 110 | 89.5 | 97.28 | 104 | 102.9096 | 116.5
| Std Dev. |  | 0.89443 | 4.5105 | 22.0749 | 1.6432 | 20.8782 | 6.935 | 4.1312 | 2.9269 | 54.8686 | 9.3327 | 2.9358 | 25.1078 | 2.5033 | 2.1398 | 4.6332 | 1.9748 | 5.7184 | 6.3246 | 3.6813 | 13.3679 | 6.8281 | 21.7662
| Rel.Std.Dev. |  | 0.8438% | 4.557% | 19.89% | 1.796% | 19.6% | 6.77% | 4.201% | 2.922% | 33.49% | 9.017% | 2.964% | 20.41% | 2.487% | 2.086% | 4.385% | 2.005% | 5.222% | 6.95% | 3.782% | 12.24% | 6.436% | 17.25%
| PDM3 |  | 1.777% | -4.959% | 6.577% | -12.15% | 2.257% | -1.648% | -5.585% | -3.824% | 57.31% | -0.6238% | -4.912% | 18.1% | -3.344% | -1.525% | 1.457% | -5.425% | 5.137% | -12.63% | -6.543% | 4.849% | 1.867% | 21.14%
| Z-Score (Absolute) |  | 0.47 | 0.64 | 1.25 | 1.81 | 0.55 | 0.09 | 0.74 | 0.45 | 9.56 | 0.07 | 0.63 | 3.14 | 0.37 | 0.07 | 0.41 | 0.71 | 1.02 | 1.89 | 0.9 | 0.97 | 0.48 | 3.64
| Analytical results for Zr in OREAS 926b (Certified Value 83 ppm) | [23 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 0
1
| 1 | 1 | 87.5162 | 89.64 | 85.4 | 80.5 | 83 | 80.1 | 86.3 | 93 | 73 | 76 | 89.3 | 74.9 | 81.2 | 89.74 | 82 | 82.7 | 85 | 87.7 | 81.514 | 77.3945 | 108 |  | 1
| 1 | 2 | 84.9585 | 86.55 | 84 | 79.3 | 79.9 | 78.3 | 84.9 | 88 | 77.4 | 75 | 88.7 | 77.4 | 81.3 | 93.07 | 83 | 79.4 | 80 | 84 | 78.327 | 82.9049 | 101 |  | 9
| 1 | 3 | 86.6182 | 90.55 | 83.9 | 79.2 | 83.6 | 78.6 | 85.3 | 91 | 73.5 | 74 | 83.2 | 78.4 | 80.7 | 92.04 | 86 | 79.6 | 78 | 87.3 | 82.452 | 76.9606 | 103 |  | 16
| 1 | 4 | 89.1077 | 87.05 | 87.4 | 82.1 | 82.5 | 80.5 | 84.5 | 87 | 75.1 | 72 | 88.5 | 75.1 | 80.4 | 89.47 | 90 | 77 | 78 | 84.3 | 82.251 | 78.0472 | 102 |  | 83.1132
| 1 | 5 | 90.234 | 91.56 | 86.8 | 81 | 81.6 | 81.3 | 87.2 | 89 | 75.1 | 73 | 91.6 | 77.1 | 81.6 | 92.38 | 87 | 82.9 | 83 | 90.9 | 85.123 | 76.9999 | 102 |  | 132
| 1 | 6 | 85.729 | 87.56 | 87 | 82.4 | 80.7 | 81.3 | 87.4 | 94 | 74.2 | 75 | 86.9 | 76.1 | 79.6 | 90.9 | 88 | 82.4 | 82 | 93.8 | 80.523 | 79.1104 | 98
| Mean |  | 87.3606 | 88.8183 | 85.75 | 80.75 | 81.8833 | 80.0167 | 85.9333 | 90.3333 | 74.7167 | 74.1667 | 88.0333 | 76.5 | 80.8 | 91.2667 | 86 | 80.6667 | 81 | 88 | 81.6983 | 78.5696 | 102.3333
| Median |  | 87.0672 | 88.6 | 86.1 | 80.75 | 82.05 | 80.3 | 85.8 | 90 | 74.65 | 74.5 | 88.6 | 76.6 | 80.95 | 91.47 | 86.5 | 81 | 81 | 87.5 | 81.8825 | 77.7209 | 102
| Std Dev. |  | 2.0158 | 2.0516 | 1.5489 | 1.3546 | 1.4134 | 1.3029 | 1.2176 | 2.8048 | 1.5613 | 1.472 | 2.8154 | 1.3755 | 0.72938 | 1.4685 | 3.0332 | 2.3796 | 2.8284 | 3.8032 | 2.2531 | 2.2713 | 3.266
| Rel.Std.Dev. |  | 2.307% | 2.31% | 1.806% | 1.678% | 1.726% | 1.628% | 1.417% | 3.105% | 2.09% | 1.985% | 3.198% | 1.798% | 0.9027% | 1.609% | 3.527% | 2.95% | 3.492% | 4.322% | 2.758% | 2.891% | 3.192%
| PDM3 |  | 5.11% | 6.864% | 3.173% | -2.843% | -1.48% | -3.726% | 3.393% | 8.687% | -10.1% | -10.76% | 5.92% | -7.957% | -2.783% | 9.81% | 3.473% | -2.944% | -2.543% | 5.88% | -1.702% | -5.467% | 23.13%
| Z-Score (Absolute) |  | 0.9 | 1.14 | 0.63 | 0.19 | 0 | 0.31 | 0.66 | 1.38 | 1.17 | 1.26 | 1.01 | 0.88 | 0.18 | 1.54 | 0.67 | 0.2 | 0.14 | 1 | 0.03 | 0.54 | 3.35
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1055: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 926b (Certified Value 3.78 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 3
| 1 | 1 | 3.8 | 3.58 | 3.88 | 3.81 | 3.73 | 3.79 | 3.67 | 3.87 | 4.45 | 3.624 | 3.9 | 3.8 | 3.42 | 3.75 | 3.9 | 3.84 | 3.96 | 3.78 | 3.94 | 3.8201 | 4.24 | 1
| 1 | 2 | 3.9 | 3.53 | 4.09 | 3.79 | 3.69 | 3.72 | 3.75 | 3.97 | 4.1 | 3.718 | 3.8 | 3.89 | 3.57 | 3.75 | 4 | 3.75 | 4.49 | 3.6 | 3.63 | 3.7088 | 3.87 | 11
| 1 | 3 | 3.9 | 3.66 | 3.94 | 3.77 | 3.87 | 3.61 | 3.68 | 3.62 | 3.98 | 3.664 | 3.8 | 3.86 | 3.44 | 3.7 | 3.9 | 3.65 | 3.84 | 4.1 | 3.43 | 3.7384 | 4.03 | 16
| 1 | 4 | 3.8 | 3.79 | 3.99 | 3.82 | 3.92 | 3.62 | 3.73 | 3.84 | 4.05 | 3.617 | 3.8 | 3.77 | 3.33 | 3.7 | 4 | 3.58 | 4.35 | 4.25 | 3.58 | 3.7147 | 3.91 | 3.7834
| 1 | 5 | 5 | 3.53 | 4.05 | 3.8 | 3.8 | 3.73 | 3.71 | 3.74 | 4.04 | 3.58 | 3.8 | 3.74 | 3.71 | 3.8 | 4 | 3.61 | 4.39 | 4.04 | 3.51 | 3.7376 | 4.16 | 134
| 1 | 6 | 4 | 3.49 | 4.18 | 3.8 | 3.74 | 3.67 | 3.7 | 3.4 | 3.83 | 3.639 | 3.9 | 3.77 | 3.46 | 3.8 | 3.9 | 3.52 | 4.49 | 3.21 | 3.54 | 3.714 | 3.91
| Mean |  | 4.0667 | 3.5967 | 4.0217 | 3.7983 | 3.7917 | 3.69 | 3.7067 | 3.74 | 4.075 | 3.6403 | 3.8333 | 3.805 | 3.4883 | 3.75 | 3.95 | 3.6583 | 4.2533 | 3.83 | 3.605 | 3.7389 | 4.02
| Median |  | 3.9 | 3.555 | 4.02 | 3.8 | 3.77 | 3.695 | 3.705 | 3.79 | 4.045 | 3.6315 | 3.8 | 3.785 | 3.45 | 3.75 | 3.95 | 3.63 | 4.37 | 3.91 | 3.56 | 3.7262 | 3.97
| Std Dev. |  | 0.46332 | 0.1113 | 0.10797 | 0.017224 | 0.088863 | 0.06957 | 0.030111 | 0.20484 | 0.20599 | 0.047001 | 0.05164 | 0.058224 | 0.13318 | 0.044721 | 0.054772 | 0.11754 | 0.28176 | 0.38294 | 0.1774 | 0.04172 | 0.15153
| Rel.Std.Dev. |  | 11.39% | 3.094% | 2.685% | 0.4535% | 2.344% | 1.885% | 0.8123% | 5.477% | 5.055% | 1.291% | 1.347% | 1.53% | 3.818% | 1.193% | 1.387% | 3.213% | 6.624% | 9.998% | 4.921% | 1.116% | 3.769%
| PDM3 |  | 7.487% | -4.936% | 6.297% | 0.3944% | 0.2181% | -2.469% | -2.029% | -1.147% | 7.707% | -3.782% | 1.319% | 0.5706% | -7.799% | -0.8832% | 4.403% | -3.306% | 12.42% | 1.231% | -4.716% | -1.176% | 6.253%
| Z-Score (Absolute) |  | 1.39 | 0.99 | 1.16 | 0.03 | 0 | 0.51 | 0.43 | 0.26 | 1.43 | 0.77 | 0.21 | 0.07 | 1.53 | 0.21 | 0.8 | 0.67 | 2.33 | 0.19 | 0.94 | 0.27 | 1.15
| Analytical results for Al in OREAS 926b (Certified Value 0.479 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  |  | 3
| 1 | 1 | 0.46 | 0.47 | 0.49 | 0.44 | 0.46 | 0.58 | 0.42 | 0.47 | 0.58 | 0.6297 | 0.5 | 0.51 | 0.49 | 0.29 | 0.62143 | 0.34 | 0.564 | 0.302 |  |  |  | 1
| 1 | 2 | 0.46 | 0.46 | 0.5 | 0.44 | 0.47 | 0.59 | 0.41 | 0.48 | 0.592 | 0.62301 | 0.5 | 0.51 | 0.49 | 0.34 | 0.62723 | 0.36 | 0.561 | 0.307 |  |  |  | #N/A
| 1 | 3 | 0.47 | 0.46 | 0.46 | 0.45 | 0.46 | 0.53 | 0.41 | 0.46 | 0.595 | 0.62584 | 0.48 | 0.52 | 0.49 | 0.3 | 0.61735 | 0.37 | 0.561 | 0.302 |  |  |  | 16
| 1 | 4 | 0.46 | 0.46 | 0.46 | 0.46 | 0.46 | 0.57 | 0.41 | 0.43 | 0.581 | 0.62263 | 0.49 | 0.5 | 0.5 | 0.34 | 0.62301 | 0.37 | 0.568 | 0.307 |  |  |  | 0.47935
| 1 | 5 | 0.46 | 0.46 | 0.46 | 0.45 | 0.44 | 0.57 | 0.42 | 0.43 | 0.608 | 0.62201 | 0.48 | 0.52 | 0.5 | 0.35 | 0.62891 | 0.36 | 0.561 | 0.307 |  |  |  | 135
| 1 | 6 | 0.47 | 0.47 | 0.44 | 0.45 | 0.44 | 0.58 | 0.39 | 0.47 | 0.593 | 0.6208 | 0.49 | 0.51 | 0.51 | 0.35 | 0.61258 | 0.36 | 0.581 | 0.296
| Mean |  | 0.46333 | 0.46333 | 0.46833 | 0.44833 | 0.455 | 0.57 | 0.41 | 0.45667 | 0.5915 | 0.624 | 0.49 | 0.51167 | 0.49667 | 0.32833 | 0.62175 | 0.36 | 0.566 | 0.3035
| Median |  | 0.46 | 0.46 | 0.46 | 0.45 | 0.46 | 0.575 | 0.41 | 0.465 | 0.5925 | 0.62282 | 0.49 | 0.51 | 0.495 | 0.34 | 0.62222 | 0.36 | 0.5625 | 0.3045
| Std Dev. |  | 0.005164 | 0.005164 | 0.022286 | 0.0075277 | 0.012247 | 0.020976 | 0.010954 | 0.021602 | 0.010291 | 0.003256 | 0.0089443 | 0.0075277 | 0.008165 | 0.026394 | 0.0061059 | 0.010954 | 0.0078486 | 0.0044159
| Rel.Std.Dev. |  | 1.115% | 1.115% | 4.759% | 1.679% | 2.692% | 3.68% | 2.672% | 4.73% | 1.74% | 0.5218% | 1.825% | 1.471% | 1.644% | 8.039% | 0.982% | 3.043% | 1.387% | 1.455%
| PDM3 |  | -3.342% | -3.342% | -2.299% | -6.471% | -5.08% | 18.91% | -14.47% | -4.733% | 23.4% | 30.18% | 2.221% | 6.741% | 3.612% | -31.5% | 29.71% | -24.9% | 18.08% | -36.69%
| Z-Score (Absolute) |  | 0.03 | 0.03 | 0.03 | 0.23 | 0.14 | 1.38 | 0.74 | 0.12 | 1.67 | 2.1 | 0.32 | 0.61 | 0.41 | 1.82 | 2.07 | 1.4 | 1.33 | 2.15
| Analytical results for As in OREAS 926b (Certified Value 96 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  | 1
| 1 | 1 | 96.25 | 92.3 | 100.5 | 98.6 | 97.8 | 95.4 | 94.5 | 104 | 98.1 | 95.4 | 97 | 95.8 | 95 | 96 | 97 | 98 | 95.893 | 88.4 | 90.7035 | 86 |  | 1
| 1 | 2 | 96.32 | 99.2 | 101 | 99.2 | 95.2 | 97.5 | 99.1 | 103 | 99.4 | 94.9 | 96 | 95.49 | 95 | 96 | 98 | 95 | 98.298 | 89.2 | 89.4416 | 88 |  | 1
| 1 | 3 | 94.62 | 94.2 | 99.6 | 101 | 96.5 | 96 | 89.9 | 98.5 | 96.1 | 96.5 | 95 | 95.57 | 93 | 95 | 95 | 94 | 100.294 | 89.5 | 91.3072 | 91 |  | 16
| 1 | 4 | 95 | 96 | 100.5 | 104.5 | 97.1 | 98.6 | 95.7 | 101 | 99.5 | 96.4 | 97 | 95.35 | 91.5 | 96 | 88 | 92 | 101.151 | 90.1 | 92.5524 | 87 |  | 96.4435
| 1 | 5 | 96.01 | 94.9 | 99.7 | 99.1 | 99.4 | 96 | 91.4 | 104 | 96.1 | 94.4 | 98 | 95.43 | 96 | 96 | 92 | 94 | 104.353 | 86.1 | 94.0619 | 85 |  | 136
| 1 | 6 | 94.49 | 98 | 99.7 | 99.5 | 98.6 | 96 | 99.4 | 97.5 | 97.4 | 96 | 97 | 95.68 | 91.5 | 96 | 89 | 99 | 97.318 | 87.4 | 91.7331 | 88
| Mean |  | 95.4483 | 95.7667 | 100.1667 | 100.3167 | 97.4333 | 96.5833 | 95 | 101.3333 | 97.7667 | 95.6 | 96.6667 | 95.5533 | 93.6667 | 95.8333 | 93.1667 | 95.3333 | 99.5512 | 88.45 | 91.6333 | 87.5
| Median |  | 95.505 | 95.45 | 100.1 | 99.35 | 97.45 | 96 | 95.1 | 102 | 97.75 | 95.7 | 97 | 95.53 | 94 | 96 | 93.5 | 94.5 | 99.296 | 88.8 | 91.5201 | 87.5
| Std Dev. |  | 0.83946 | 2.5319 | 0.5785 | 2.2049 | 1.5055 | 1.2106 | 3.8956 | 2.8225 | 1.5148 | 0.84617 | 1.0328 | 0.16597 | 1.9408 | 0.40825 | 4.1673 | 2.6583 | 3.0367 | 1.4843 | 1.5829 | 2.0736
| Rel.Std.Dev. |  | 0.8795% | 2.644% | 0.5775% | 2.198% | 1.545% | 1.253% | 4.101% | 2.785% | 1.549% | 0.8851% | 1.068% | 0.1737% | 2.072% | 0.426% | 4.473% | 2.788% | 3.05% | 1.678% | 1.727% | 2.37%
| PDM3 |  | -1.032% | -0.7018% | 3.86% | 4.016% | 1.026% | 0.145% | -1.497% | 5.07% | 1.372% | -0.8746% | 0.2314% | -0.923% | -2.879% | -0.6327% | -3.398% | -1.151% | 3.222% | -8.288% | -4.988% | -9.273%
| Z-Score (Absolute) |  | 0.08 | 0.03 | 1.61 | 1.66 | 0.63 | 0.32 | 0.24 | 2.02 | 0.75 | 0.03 | 0.35 | 0.05 | 0.72 | 0.05 | 0.9 | 0.13 | 1.38 | 2.59 | 1.45 | 2.93
| Analytical results for Au in OREAS 926b (Indicative Value 0.017 ppm) | [22 empty] | Indicative
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 11 | 15 | 20 | 22 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 40g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | 0.0116 | < 0.02 | < 0.02 | 0.02 | < 0.02 | 0.0024 | 0.0181 | 0.11 | 0.029 | 0.03 | [11 empty] | 1
| 1 | 2 | 0.0116 | < 0.02 | < 0.02 | 0.02 | 0.02 | 0.0021 | 0.0181 | 0.115 | 0.031 | 0.04 | [11 empty] | 19
| 1 | 3 | 0.0115 | < 0.02 | < 0.02 | 0.02 | 0.02 | 0.0025 | 0.0163 | 0.118 | 0.089 | 0.02 | [11 empty] | 16
| 1 | 4 | 0.0116 | < 0.02 | < 0.02 | 0.02 | < 0.02 | 0.001 | 0.0175 | 0.107 | 0.071 | 0.03 | [11 empty] | 0.017214
| 1 | 5 | 0.0109 | 0.02 | < 0.02 | 0.02 | < 0.02 | 0.001 | 0.0179 | 0.104 | 0.062 | 0.03 | [11 empty] | 25
| 1 | 6 | 0.0108 | 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.0014 | 0.0167 | 0.105 | 0.055 | 0.03
| Mean |  | 0.011333 | 0.02 |  | 0.02 | 0.02 | 0.0017333 | 0.017433 | 0.10983 | 0.056167 | 0.03
| Median |  | 0.01155 | 0.02 |  | 0.02 | 0.02 | 0.00175 | 0.0177 | 0.1085 | 0.0585 | 0.03
| Std Dev. |  | 0.00037771 | 0 |  | 0 | 0 | 0.00068605 | 0.00076594 | 0.0056362 | 0.023259 | 0.0063246
| Rel.Std.Dev. |  | 3.333% | 0% |  | 0% | 0% | 39.58% | 4.394% | 5.132% | 41.41% | 21.08%
| PDM3 |  | -34.16% | 16.18% |  | 16.18% | 16.18% | -89.93% | 1.272% | 538% | 226.3% | 74.27%
| Z-Score (Absolute) |  | 0.58 | 0.29 | 0.77 | 0.42 | 0.29 | 1.95 | 0.29 | 13.5 | 5.83 | 2.09
| Analytical results for B in OREAS 926b (Certified Value < 10 ppm) | [22 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 11 | 17 | 18 | 20 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [11 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | [11 empty] | 1
| 1 | 1 | < 10 | < 10 | < 10 | < 10 | < 10 | 4 | < 20 | < 10 | < 10 | < 5 | [11 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | < 10 | < 5 | [11 empty] | 16
| 1 | 3 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | < 10 | < 5 | [11 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | 4 | < 20 | < 10 | < 10 | < 5 | [11 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | < 10 | < 5 | [11 empty] | 137
| 1 | 6 | < 10 | < 10 | < 10 | < 10 | < 10 | 3 | < 20 | < 10 | < 10 | < 5
| Mean | [6 empty] | 3.3333
| Median | [6 empty] | 3
| Std Dev. | [6 empty] | 0.5164
| Rel.Std.Dev. | [6 empty] | 15.49%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ba in OREAS 926b (Indicative Value 105 ppm) | [22 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 1761.79 | 106.5 | 110 | 120 | 110 | 110 | 35.7 | 70 | 111.4 | 152 | 209.17 | 82 | 128 | 115 | 61 | 146.881 | 73 | 136.3708 | 57 |  |  | 1
| 1 | 2 | 1748.05 | 110 | 100 | 120 | 110 | 110 | 19.6 | 70 | 112.9 | 151 | 206.72 | 67 | 131 | 119 | 72 | 145.086 | 74 | 133.6955 | 51 |  |  | 20
| 1 | 3 | 1730.52 | 108.5 | 100 | 110 | 110 | 110 | 26.4 | 60 | 104.9 | 148 | 207.6 | 85 | 129 | 117 | 62 | 146.209 | 76 | 134.392 | 58 |  |  | 16
| 1 | 4 | 1754.51 | 108.5 | 110 | 110 | 110 | 110 | 18.1 | 60 | 98.3 | 150 | 205.93 | 77 | 129 | 113 | 70 | 150.304 | 77 | 137.9828 | 53 |  |  | 105.3058
| 1 | 5 | 1761.83 | 109 | 100 | 110 | 110 | 100 | 22 | 60 | 98.5 | 159 | 206.86 | 76 | 133 | 115 | 73 | 148.338 | 74 | 134.8184 | 56 |  |  | 26
| 1 | 6 | 1742.77 | 111 | 110 | 100 | 110 | 100 | 20.6 | 60 | 109.4 | 161 | 202 | 68 | 131 | 116 | 74 | 147.595 | 73 | 135.1369 | 55
| Mean |  | 1749.9117 | 108.9167 | 105 | 111.6667 | 110 | 106.6667 | 23.7333 | 63.3333 | 105.9 | 153.5 | 206.38 | 75.8333 | 130.1667 | 115.8333 | 68.6667 | 147.4022 | 74.5 | 135.3994 | 55
| Median |  | 1751.28 | 108.75 | 105 | 110 | 110 | 110 | 21.3 | 60 | 107.15 | 151.5 | 206.79 | 76.5 | 130 | 115.5 | 71 | 147.238 | 74 | 134.9777 | 55.5
| Std Dev. |  | 12.1226 | 1.5303 | 5.4772 | 7.5277 | 0 | 5.164 | 6.5108 | 5.164 | 6.4034 | 5.244 | 2.4085 | 7.2503 | 1.8348 | 2.0412 | 5.7155 | 1.8107 | 1.6432 | 1.5454 | 2.6077
| Rel.Std.Dev. |  | 0.6928% | 1.405% | 5.216% | 6.741% | 0% | 4.841% | 27.43% | 8.154% | 6.047% | 3.416% | 1.167% | 9.561% | 1.41% | 1.762% | 8.324% | 1.228% | 2.206% | 1.141% | 4.741%
| PDM3 |  | 1562% | 3.429% | -0.2904% | 6.04% | 4.458% | 1.292% | -77.46% | -39.86% | 0.5642% | 45.77% | 95.98% | -27.99% | 23.61% | 9.997% | -34.79% | 39.98% | -29.25% | 28.58% | -47.77%
| Z-Score (Absolute) |  | NA | 0.03 | 0.06 | 0.09 | 0.05 | 0.03 | 1.9 | 1.01 | 0.04 | 1.03 | 2.23 | 0.72 | 0.51 | 0.18 | 0.89 | 0.9 | 0.75 | 0.63 | 1.2
| Analytical results for Be in OREAS 926b (Certified Value 0.31 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 17 | 18 | 19 | 20 | 21 | 22 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 0.33 | 0.35 | 0.36 | 0.33 | 0.34 | 0.5 | 0.09 | < 0.5 | 0.31 | 0.2 | < 5 | 0.34 | 0.47701 | 0.24 | [7 empty] | 1
| 1 | 2 | 0.35 | 0.34 | 0.36 | 0.34 | 0.34 | 0.4 | 0.28 | < 0.5 | 0.32 | 0.2 | < 5 | 0.31 | 0.47752 | 0.25 | [7 empty] | #N/A
| 1 | 3 | 0.35 | 0.34 | 0.35 | 0.33 | 0.34 | 0.4 | 0.16 | < 0.5 | 0.3 | 0.2 | < 5 | 0.22 | 0.52696 | 0.22 | [7 empty] | 16
| 1 | 4 | 0.34 | 0.34 | 0.35 | 0.33 | 0.33 | 0.5 | 0.25 | < 0.5 | 0.3 | 0.2 | < 5 | 0.31 | 0.49369 | 0.2 | [7 empty] | 0.31271
| 1 | 5 | 0.34 | 0.33 | 0.34 | 0.33 | 0.32 | 0.4 | 0.17 | < 0.5 | 0.29 | 0.3 | < 5 | 0.27 | 0.48048 | 0.2 | [7 empty] | 138
| 1 | 6 | 0.35 | 0.34 | 0.34 | 0.33 | 0.32 | 0.4 | 0.26 | < 0.5 | 0.31 | 0.3 | < 5 | 0.22 | 0.49144 | 0.22
| Mean |  | 0.34333 | 0.34 | 0.35 | 0.33167 | 0.33167 | 0.43333 | 0.20167 |  | 0.305 | 0.23333 |  | 0.27833 | 0.49118 | 0.22167
| Median |  | 0.345 | 0.34 | 0.35 | 0.33 | 0.335 | 0.4 | 0.21 |  | 0.305 | 0.2 |  | 0.29 | 0.48596 | 0.22
| Std Dev. |  | 0.008165 | 0.0063246 | 0.0089443 | 0.0040825 | 0.0098319 | 0.05164 | 0.073598 |  | 0.010488 | 0.05164 |  | 0.050365 | 0.018915 | 0.020412
| Rel.Std.Dev. |  | 2.378% | 1.86% | 2.556% | 1.231% | 2.964% | 11.92% | 36.49% |  | 3.439% | 22.13% |  | 18.1% | 3.851% | 9.209%
| PDM3 |  | 9.793% | 8.728% | 11.93% | 6.063% | 6.063% | 38.57% | -35.51% |  | -2.465% | -25.38% |  | -10.99% | 57.07% | -29.11%
| Z-Score (Absolute) |  | 0.2 | 0.14 | 0.31 | 0 | 0 | NA | 2.19 | 1.38 | 0.45 | NA | 36.55 | 0.9 | 2.69 | 1.85
Results from laboratories 8 and 18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 926b (Certified Value 47.8 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 48.6 | 45.7 | 49.6 | 52.9 | 47.4 | 47.3 | 51 | 47.06 | 52.75 | 47.5 | 48.8 | 54.4 | 45 | 52.53 | 50.641 | 49.4 | 43.9505 | 51.3 |  |  |  | 1
| 1 | 2 | 51.1 | 46.2 | 52.4 | 51.2 | 49.6 | 47.9 | 48.6 | 46.88 | 54.08 | 47.9 | 49 | 55.9 | 46 | 58.39 | 52.146 | 48.9 | 43.0143 | 50.4 |  |  |  | 2
| 1 | 3 | 48.8 | 45.8 | 48.2 | 45 | 46 | 45.4 | 48.9 | 46.19 | 55.64 | 46.8 | 48.6 | 54.8 | 45 | 51 | 48.839 | 48.6 | 44.0425 | 51.2 |  |  |  | 16
| 1 | 4 | 50.4 | 46.4 | 49.2 | 45.5 | 48.1 | 48.6 | 50.2 | 46.18 | 55.96 | 47.2 | 49.7 | 55.3 | 46 | 56.55 | 47.448 | 49.3 | 44.8585 | 50.5 |  |  |  | 47.8303
| 1 | 5 | 48.4 | 45.6 | 48.1 | 46 | 49.4 | 47.3 | 50.2 | 45.52 | 55.21 | 47.8 | 49.9 | 56.6 | 45 | 56.49 | 42.642 | 48.2 | 45.1191 | 51.5 |  |  |  | 139
| 1 | 6 | 51.9 | 46.9 | 46.7 | 45.9 | 49 | 49.6 | 48.2 | 44.96 | 53.02 | 46.4 | 46.9 | 55.9 | 45 | 58.27 | 46.091 | 47.8 | 44.141 | 50.2
| Mean |  | 49.8667 | 46.1 | 49.0333 | 47.75 | 48.25 | 47.6833 | 49.5167 | 46.1317 | 54.4433 | 47.2667 | 48.8167 | 55.4833 | 45.3333 | 55.5383 | 47.9678 | 48.7 | 44.1877 | 50.85
| Median |  | 49.6 | 46 | 48.7 | 45.95 | 48.55 | 47.6 | 49.55 | 46.185 | 54.645 | 47.35 | 48.9 | 55.6 | 45 | 56.52 | 48.1435 | 48.75 | 44.0917 | 50.85
| Std Dev. |  | 1.472 | 0.498 | 1.9336 | 3.3922 | 1.3795 | 1.419 | 1.1035 | 0.7968 | 1.3671 | 0.58538 | 1.0685 | 0.8085 | 0.5164 | 3.0715 | 3.3917 | 0.6261 | 0.74486 | 0.54681
| Rel.Std.Dev. |  | 2.952% | 1.08% | 3.943% | 7.104% | 2.859% | 2.976% | 2.229% | 1.727% | 2.511% | 1.238% | 2.189% | 1.457% | 1.139% | 5.53% | 7.071% | 1.286% | 1.686% | 1.075%
| PDM3 |  | 4.258% | -3.617% | 2.515% | -0.1678% | 0.8776% | -0.3072% | 3.526% | -3.551% | 13.83% | -1.178% | 2.062% | 16% | -5.22% | 16.12% | 0.2876% | 1.818% | -7.616% | 6.313%
| Z-Score (Absolute) |  | 0.72 | 1.23 | 0.29 | 0.38 | 0.12 | 0.41 | 0.54 | 1.22 | 3.1 | 0.63 | 0.18 | 3.64 | 1.63 | 3.66 | 0.26 | 0.12 | 2.22 | 1.23
| Analytical results for Ca in OREAS 926b (Certified Value 0.03 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  |  | 3
| 1 | 1 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0486 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.027701 | 0.04 | 0.0324 | 0.029 |  |  |  | 1
| 1 | 2 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0486 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027561 | 0.04 | 0.0324 | 0.029 |  |  |  | 12
| 1 | 3 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0484 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.028569 | 0.04 | 0.0324 | 0.029 |  |  |  | 16
| 1 | 4 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.049 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.028402 | 0.04 | 0.0333 | 0.029 |  |  |  | 0.030293
| 1 | 5 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0487 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.027868 | 0.04 | 0.0405 | 0.029 |  |  |  | 140
| 1 | 6 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0487 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027299 | 0.04 | 0.04 | 0.029
| Mean |  | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.048667 | 0.03 | 0.03 | 0.03 | 0.03 | 0.023333 | 0.0279 | 0.04 | 0.035167 | 0.029
| Median |  | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04865 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.027785 | 0.04 | 0.03285 | 0.029
| Std Dev. |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00019664 | 0 | 0 | 0 | 0 | 0.005164 | 0.00049329 | 0 | 0.0039561 | 0
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0.4041% | 0% | 0% | 0% | 0% | 22.13% | 1.768% | 0% | 11.25% | 0%
| PDM3 |  | -0.9685% | -0.9685% | -0.9685% | -0.9685% | -0.9685% | -0.9685% | -0.9685% | -0.9685% | 60.65% | -0.9685% | -0.9685% | -0.9685% | -0.9685% | -22.98% | -7.901% | 32.04% | 16.09% | -4.27%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cd in OREAS 926b (Certified Value 0.15 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.203 | 0.14 | 0.15 | 0.15 | 0.18 | 0.13 | 3.71 | 0.14 | 0.2 | 0.2 | < 0.5 | 0.19 | 0.21 | < 5 | 0.19 | 0.16903 | 0.73 |  |  |  |  | 1
| 1 | 2 | 0.149 | 0.15 | 0.14 | 0.14 | 0.15 | 0.19 | 3.22 | 0.11 | 0.2 | 0.3 | < 0.5 | 0.19 | 0.15 | < 5 | 0.14 | 0.16852 | 0.28 |  |  |  |  | #N/A
| 1 | 3 | 0.149 | 0.13 | 0.13 | 0.16 | 0.17 | 0.13 | 3.09 | 0.21 | 0.19 | 0.2 | < 0.5 | 0.17 | 0.12 | < 5 | 0.17 | 0.16299 | 0.22 |  |  |  |  | 16
| 1 | 4 | 0.157 | 0.14 | 0.14 | 0.15 | 0.16 | 0.12 | 3.07 | 0.1 | 0.19 | 0.2 | < 0.5 | 0.17 | 0.14 | < 5 | 0.16 | 0.1614 | 0.23 |  |  |  |  | 0.15479
| 1 | 5 | 0.141 | 0.14 | 0.13 | 0.13 | 0.15 | 0.15 | 3.1 | 0.06 | 0.19 | 0.2 | < 0.5 | 0.17 | 0.14 | < 5 | 0.14 | 0.16003 | 0.59 |  |  |  |  | 141
| 1 | 6 | 0.148 | 0.15 | 0.15 | 0.14 | 0.14 | 0.14 | 3.09 | 0.08 | 0.2 | 0.2 | < 0.5 | 0.16 | 0.14 | < 5 | 0.15 | 0.16774 | 0.24
| Mean |  | 0.15783 | 0.14167 | 0.14 | 0.145 | 0.15833 | 0.14333 | 3.2133 | 0.11667 | 0.195 | 0.21667 |  | 0.175 | 0.15 |  | 0.15833 | 0.16495 | 0.38167
| Median |  | 0.149 | 0.14 | 0.14 | 0.145 | 0.155 | 0.135 | 3.095 | 0.105 | 0.195 | 0.2 |  | 0.17 | 0.14 |  | 0.155 | 0.16536 | 0.26
| Std Dev. |  | 0.022702 | 0.0075277 | 0.0089443 | 0.010488 | 0.01472 | 0.025033 | 0.24921 | 0.053166 | 0.0054772 | 0.040825 |  | 0.012247 | 0.030984 |  | 0.019408 | 0.0039428 | 0.22104
| Rel.Std.Dev. |  | 14.38% | 5.314% | 6.389% | 7.233% | 9.297% | 17.47% | 7.756% | 45.57% | 2.809% | 18.84% |  | 6.999% | 20.66% |  | 12.26% | 2.39% | 57.91%
| PDM3 |  | 1.968% | -8.477% | -9.553% | -6.323% | 2.291% | -7.4% | 1976% | -24.63% | 25.98% | 39.98% |  | 13.06% | -3.093% |  | 2.291% | 6.566% | 146.6%
| Z-Score (Absolute) |  | 0.02 | 0.67 | 0.74 | 0.54 | 0 | 0.61 | 123.6 | 1.69 | 1.48 | NA | 3.71 | 0.67 | 0.34 | 94.74 | 0 | 0.27 | 9.04
| Analytical results for Ce in OREAS 926b (Certified Value 49.3 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 48.2 | 48.8 | 55.4 | 48.7 | 51.1 | 58.9 | 43 | 57 | 64.8 | 49.8 | 32 | 59.645 | 42.5 | 48.7756 | 32 | [6 empty] | 1
| 1 | 2 | 50.2 | 47.2 | 55.9 | 48.1 | 52.8 | 61.9 | 41.3 | 57.1 | 64.8 | 51 | 38.14 | 56.515 | 43.4 | 47.7478 | 33 | [6 empty] | #N/A
| 1 | 3 | 49.1 | 47.3 | 51.8 | 48.5 | 51.4 | 55 | 41.6 | 56 | 64.4 | 50 | 34.57 | 57.723 | 43.6 | 48.1284 | 32 | [6 empty] | 16
| 1 | 4 | 50.4 | 47.7 | 52.1 | 46.6 | 51.4 | 58.1 | 42 | 56.3 | 66 | 50 | 38.08 | 58.828 | 43.4 | 48.9045 | 33 | [6 empty] | 49.2938
| 1 | 5 | 49.7 | 46.5 | 49.1 | 49.4 | 50.6 | 56.9 | 42.1 | 57.3 | 66.2 | 49.1 | 38.81 | 58.607 | 42.9 | 48.6879 | 32 | [6 empty] | 142
| 1 | 6 | 50.6 | 48.5 | 48.3 | 48 | 51.3 | 63 | 40.4 | 56.1 | 61.4 | 49.1 | 38.46 | 57.055 | 43.3 | 48.3623 | 31
| Mean |  | 49.7 | 47.6667 | 52.1 | 48.2167 | 51.4333 | 58.9667 | 41.7333 | 56.6333 | 64.6 | 49.8333 | 36.6767 | 58.0622 | 43.1833 | 48.4344 | 32.1667
| Median |  | 49.95 | 47.5 | 51.95 | 48.3 | 51.35 | 58.5 | 41.8 | 56.65 | 64.8 | 49.9 | 38.11 | 58.165 | 43.35 | 48.5251 | 32
| Std Dev. |  | 0.91214 | 0.85946 | 3.126 | 0.93684 | 0.73394 | 3.021 | 0.87101 | 0.56451 | 1.7251 | 0.70616 | 2.7623 | 1.1757 | 0.40702 | 0.44089 | 0.75277
| Rel.Std.Dev. |  | 1.835% | 1.803% | 6% | 1.943% | 1.427% | 5.123% | 2.087% | 0.9968% | 2.67% | 1.417% | 7.531% | 2.025% | 0.9425% | 0.9103% | 2.34%
| PDM3 |  | 0.8241% | -3.301% | 5.693% | -2.185% | 4.34% | 19.62% | -15.34% | 14.89% | 31.05% | 1.095% | -25.6% | 17.79% | -12.4% | -1.743% | -34.74%
| Z-Score (Absolute) |  | 0 | 0.21 | 0.25 | 0.15 | 0.18 | 0.96 | 0.82 | 0.72 | 1.54 | 0.01 | 1.35 | 0.87 | 0.67 | 0.13 | 1.81
| Analytical results for Co in OREAS 926b (Certified Value 6.78 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 6.4 | 6.8 | 7 | 6.8 | 6.7 | 6.8 | 6.8 | 6.8 | 6.7 | 6.7914 | 8 | 7.2 | 7.1 | 6.1 | 6.411 | 6.4 | 6.6843 | 6.92 |  |  |  | 1
| 1 | 2 | 6.79 | 6.9 | 7.1 | 6.9 | 6.8 | 7.2 | 6.4 | 7 | 6.8 | 6.674 | 7 | 7.4 | 7 | 6.8 | 6.367 | 6.5 | 6.616 | 6.71 |  |  |  | 4
| 1 | 3 | 6.58 | 6.8 | 7.2 | 7 | 6.7 | 6.2 | 6.3 | 7 | 6.7 | 6.8456 | 7 | 7.4 | 7.3 | 5.8 | 6.669 | 6.6 | 6.7208 | 7.15 |  |  |  | 16
| 1 | 4 | 6.66 | 6.9 | 6.7 | 7.5 | 6.9 | 6.9 | 6.9 | 7.1 | 6.7 | 6.3612 | 7 | 7.3 | 7 | 6.5 | 6.843 | 6.5 | 6.6277 | 6.92 |  |  |  | 6.7782
| 1 | 5 | 6.61 | 6.8 | 6.9 | 7.5 | 6.7 | 6.5 | 6.3 | 6.9 | 6.8 | 6.3026 | 7 | 7.2 | 7 | 6.5 | 6.128 | 6.4 | 6.7114 | 7.08 |  |  |  | 143
| 1 | 6 | 6.78 | 6.9 | 6.4 | 7.2 | 6.6 | 6.9 | 6.4 | 6.9 | 6.6 | 6.2882 | 7 | 7.3 | 7.1 | 6.6 | 6.747 | 6.4 | 6.6462 | 6.94
| Mean |  | 6.6367 | 6.85 | 6.8833 | 7.15 | 6.7333 | 6.75 | 6.5167 | 6.95 | 6.7167 | 6.5438 | 7.1667 | 7.3 | 7.0833 | 6.3833 | 6.5275 | 6.4667 | 6.6677 | 6.9533
| Median |  | 6.635 | 6.85 | 6.95 | 7.1 | 6.7 | 6.85 | 6.4 | 6.95 | 6.7 | 6.5176 | 7 | 7.3 | 7.05 | 6.5 | 6.54 | 6.45 | 6.6653 | 6.93
| Std Dev. |  | 0.14459 | 0.054772 | 0.29269 | 0.30166 | 0.10328 | 0.35071 | 0.26394 | 0.10488 | 0.075277 | 0.25541 | 0.40825 | 0.089443 | 0.1169 | 0.3656 | 0.2708 | 0.08165 | 0.044158 | 0.15253
| Rel.Std.Dev. |  | 2.179% | 0.7996% | 4.252% | 4.219% | 1.534% | 5.196% | 4.05% | 1.509% | 1.121% | 3.903% | 5.696% | 1.225% | 1.65% | 5.727% | 4.149% | 1.263% | 0.6623% | 2.194%
| PDM3 |  | -2.088% | 1.06% | 1.551% | 5.486% | -0.6616% | -0.4158% | -3.858% | 2.535% | -0.9075% | -3.457% | 5.731% | 7.699% | 4.502% | -5.825% | -3.698% | -4.596% | -1.629% | 2.584%
| Z-Score (Absolute) |  | 0.32 | 0.38 | 0.49 | 1.36 | 0 | 0.05 | 0.71 | 0.71 | 0.05 | 0.62 | NA | 1.86 | 1.15 | 1.15 | 0.67 | 0.87 | 0.21 | 0.72
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 926b (Certified Value 21.9 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 20.7 | 22 | 22 | 21 | 21 | 24 | 28 | 22.2 | 22 | 20 | 23.1 | 23 | 18 | 25.338 | 20.9 | 20.7196 | 20 |  |  |  |  | 1
| 1 | 2 | 21.4 | 21 | 22 | 21 | 21 | 25 | 27 | 22.3 | 22 | 20 | 23.1 | 23 | 21 | 24.874 | 21.5 | 22.9153 | 20 |  |  |  |  | 5
| 1 | 3 | 20.5 | 21 | 22 | 21 | 21 | 22 | 26 | 22 | 22 | 20 | 23.5 | 22 | 17 | 25.307 | 22 | 26.3548 | 20 |  |  |  |  | 16
| 1 | 4 | 20.7 | 21 | 22 | 21 | 20 | 24 | 26 | 21.8 | 22 | 20 | 23.3 | 22 | 20 | 25.647 | 21.9 | 21.5617 | 20 |  |  |  |  | 21.8831
| 1 | 5 | 20.8 | 21 | 22 | 21 | 21 | 22 | 26 | 21.3 | 22 | 20 | 23.6 | 22 | 20 | 24.588 | 21.5 | 23.6224 | 20 |  |  |  |  | 144
| 1 | 6 | 22 | 21 | 21 | 21 | 21 | 24 | 26 | 21.7 | 22 | 20 | 23.4 | 22 | 20 | 25.371 | 21.2 | 23.1596 | 20
| Mean |  | 21.0167 | 21.1667 | 21.8333 | 21 | 20.8333 | 23.5 | 26.5 | 21.8833 | 22 | 20 | 23.3333 | 22.3333 | 19.3333 | 25.1875 | 21.5 | 23.0556 | 20
| Median |  | 20.75 | 21 | 22 | 21 | 21 | 24 | 26 | 21.9 | 22 | 20 | 23.35 | 22 | 20 | 25.3225 | 21.5 | 23.0375 | 20
| Std Dev. |  | 0.57067 | 0.40825 | 0.40825 | 0 | 0.40825 | 1.2247 | 0.83666 | 0.3656 | 0 | 0 | 0.20656 | 0.5164 | 1.5055 | 0.38456 | 0.41473 | 1.9455 | 0
| Rel.Std.Dev. |  | 2.715% | 1.929% | 1.87% | 0% | 1.96% | 5.212% | 3.157% | 1.671% | 0% | 0% | 0.8853% | 2.312% | 7.787% | 1.527% | 1.929% | 8.438% | 0%
| PDM3 |  | -3.959% | -3.274% | -0.2274% | -4.036% | -4.797% | 7.389% | 21.1% | 0.0011% | 0.5342% | -8.605% | 6.627% | 2.057% | -11.65% | 15.1% | -1.751% | 5.358% | -8.605%
| Z-Score (Absolute) |  | 0.6 | 0.5 | 0.02 | 0.61 | 0.73 | 1.18 | 3.32 | 0.02 | 0.1 | NA | 1.06 | 0.34 | 1.81 | 2.38 | 0.26 | 0.86 | 1.33
Results from laboratory 15 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 926b (Certified Value 1.01 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.938 | 1.12 | 0.99 | 1.09 | 1.07 | 1.02 | 1.02 | 1.56 | 1.6102 | 1 | 0.97 | 0.84 | 1.16 | 1.1261 | 0.79 | [6 empty] | 1
| 1 | 2 | 0.961 | 1.1 | 0.99 | 1.04 | 1.12 | 0.97 | 0.96 | 1.58 | 1.6235 | 1 | 0.94 | 0.94 | 1.19 | 1.1691 | 0.82 | [6 empty] | #N/A
| 1 | 3 | 0.937 | 1.08 | 0.97 | 0.98 | 1.12 | 0.97 | 0.96 | 1.56 | 1.6007 | 1 | 0.95 | 0.85 | 1.19 | 1.2909 | 0.82 | [6 empty] | 16
| 1 | 4 | 0.901 | 1.08 | 1 | 0.95 | 1.08 | 1.04 | 0.87 | 1.54 | 1.467 | 1 | 0.94 | 0.93 | 1.24 | 1.3032 | 0.8 | [6 empty] | 1.0095
| 1 | 5 | 0.932 | 1.08 | 0.96 | 1.02 | 0.98 | 1.04 | 0.99 | 1.55 | 1.5214 | 1 | 0.94 | 0.94 | 1.18 | 1.3228 | 0.81 | [6 empty] | 145
| 1 | 6 | 0.948 | 1.12 | 0.96 | 0.96 | 1.02 | 1.05 | 0.87 | 1.57 | 1.4744 | 1 | 0.89 | 0.9 | 1.16 | 1.2315 | 0.79
| Mean |  | 0.93617 | 1.0967 | 0.97833 | 1.0067 | 1.065 | 1.015 | 0.945 | 1.56 | 1.5496 | 1 | 0.93833 | 0.9 | 1.1867 | 1.2406 | 0.805
| Median |  | 0.9375 | 1.09 | 0.98 | 1 | 1.075 | 1.03 | 0.96 | 1.56 | 1.5611 | 1 | 0.94 | 0.915 | 1.185 | 1.2612 | 0.805
| Std Dev. |  | 0.020074 | 0.019664 | 0.017224 | 0.053541 | 0.055767 | 0.036194 | 0.062209 | 0.014142 | 0.070729 | 0 | 0.026394 | 0.045166 | 0.029439 | 0.079378 | 0.013784
| Rel.Std.Dev. |  | 2.144% | 1.793% | 1.761% | 5.319% | 5.236% | 3.566% | 6.583% | 0.9065% | 4.565% | 0% | 2.813% | 5.018% | 2.481% | 6.398% | 1.712%
| PDM3 |  | -7.26% | 8.64% | -3.083% | -0.2761% | 5.503% | 0.5494% | -6.385% | 54.54% | 53.5% | -0.9365% | -7.045% | -10.84% | 17.56% | 22.9% | -20.25%
| Z-Score (Absolute) |  | 0.63 | 0.72 | 0.27 | 0.04 | 0.46 | 0.04 | 0.55 | 4.61 | 4.53 | NA | 0.61 | 0.93 | 1.48 | 1.93 | 1.73
Results from laboratory 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 926b (Certified Value 0.873 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  | 3
| 1 | 1 | 0.856 | 0.916 | 0.886 | 0.851 | 0.861 | 0.895 | 0.892 | 0.84342 | 0.8707 | 0.858 | 0.8391 | 0.903 | 0.86254 | 0.8766 | 0.8737 | 0.92342 | 0.817 | 0.91 | 0.89 |  |  | 1
| 1 | 2 | 0.889 | 0.915 | 0.877 | 0.858 | 0.883 | 0.935 | 0.867 | 0.84814 | 0.8809 | 0.866 | 0.8364 | 0.894 | 0.87426 | 0.8939 | 0.8773 | 0.94487 | 0.835 | 0.914 | 0.881 |  |  | #N/A
| 1 | 3 | 0.885 | 0.913 | 0.861 | 0.861 | 0.874 | 0.845 | 0.867 | 0.8217 | 0.8825 | 0.838 | 0.8356 | 0.89 | 0.8765 | 0.8972 | 0.8752 | 0.91776 | 0.856 | 0.909 | 0.907 |  |  | 16
| 1 | 4 | 0.876 | 0.908 | 0.862 | 0.888 | 0.862 | 0.895 | 0.875 | 0.85702 | 0.8871 | 0.852 | 0.8342 | 0.894 | 0.8749 | 0.8778 | 0.87056 | 0.93253 | 0.847 | 0.91 | 0.881 |  |  | 0.87334
| 1 | 5 | 0.865 | 0.917 | 0.856 | 0.862 | 0.858 | 0.837 | 0.883 | 0.82742 | 0.8813 | 0.855 | 0.8362 | 0.891 | 0.86232 | 0.8714 | 0.87459 | 0.91243 | 0.838 | 0.925 | 0.912 |  |  | 146
| 1 | 6 | 0.898 | 0.925 | 0.835 | 0.872 | 0.869 | 0.874 | 0.833 | 0.83757 | 0.888 | 0.836 | 0.8359 | 0.899 | 0.86817 | 0.8801 | 0.87565 | 0.91668 | 0.84 | 0.926 | 0.904
| Mean |  | 0.87817 | 0.91567 | 0.86283 | 0.86533 | 0.86783 | 0.88017 | 0.8695 | 0.83921 | 0.88175 | 0.85083 | 0.83623 | 0.89517 | 0.86978 | 0.88283 | 0.8745 | 0.92462 | 0.83883 | 0.91567 | 0.89583
| Median |  | 0.8805 | 0.9155 | 0.8615 | 0.8615 | 0.8655 | 0.8845 | 0.871 | 0.84049 | 0.8819 | 0.8535 | 0.83605 | 0.894 | 0.87122 | 0.87895 | 0.87489 | 0.92059 | 0.839 | 0.912 | 0.897
| Std Dev. |  | 0.015664 | 0.0055737 | 0.01768 | 0.013018 | 0.0094534 | 0.03629 | 0.020315 | 0.013136 | 0.0061851 | 0.011703 | 0.0016058 | 0.0049565 | 0.0063577 | 0.010308 | 0.0022726 | 0.012106 | 0.013075 | 0.0078145 | 0.013615
| Rel.Std.Dev. |  | 1.784% | 0.6087% | 2.049% | 1.504% | 1.089% | 4.123% | 2.336% | 1.565% | 0.7015% | 1.376% | 0.192% | 0.5537% | 0.731% | 1.168% | 0.2599% | 1.309% | 1.559% | 0.8534% | 1.52%
| PDM3 |  | 0.5525% | 4.846% | -1.203% | -0.9169% | -0.6307% | 0.7815% | -0.4398% | -3.908% | 0.9628% | -2.577% | -4.249% | 2.499% | -0.4076% | 1.087% | 0.1327% | 5.871% | -3.951% | 4.846% | 2.575%
| Z-Score (Absolute) |  | 0.21 | 2.38 | 0.67 | 0.53 | 0.39 | 0.33 | 0.29 | 2.04 | 0.42 | 1.37 | 2.21 | 1.19 | 0.27 | 0.48 | 0 | 2.9 | 2.06 | 2.38 | 1.23
| Analytical results for Dy in OREAS 926b (Certified Value 1.52 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 14 | 15 | 17 | 22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [16 empty] | 2
|  |  | 0.5g | 0.25g | 40g | 0.5g | 0.5g | [16 empty] | 2
| 1 | 1 | 1.7 | 1.9011 | 1.4 | 1.36 | 1.25 | [16 empty] | 1
| 1 | 2 | 1.8 | 1.836 | 1.4 | 1.32 | 1.25 | [16 empty] | #N/A
| 1 | 3 | 1.5 | 1.7768 | 1.37 | 1.38 | 1.39 | [16 empty] | 16
| 1 | 4 | 1.7 | 1.845 | 1.44 | 1.5 | 1.3 | [16 empty] | 1.5232
| 1 | 5 | 1.6 | 1.7484 | 1.4 | 1.48 | 1.33 | [16 empty] | 147
| 1 | 6 | 1.8 | 1.8199 | 1.37 | 1.47 | 1.26
| Mean |  | 1.6833 | 1.8212 | 1.3967 | 1.4183 | 1.2967
| Median |  | 1.7 | 1.8279 | 1.4 | 1.425 | 1.28
| Std Dev. |  | 0.1169 | 0.053747 | 0.02582 | 0.074409 | 0.055737
| Rel.Std.Dev. |  | 6.945% | 2.951% | 1.849% | 5.246% | 4.299%
| PDM3 |  | 10.51% | 19.56% | -8.309% | -6.887% | -14.87%
| Z-Score (Absolute) |  | 1.47 | 2.23 | 0.12 | 0 | 0.67
| Analytical results for Er in OREAS 926b (Indicative Value 0.54 ppm) | [22 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 08 | 15 | 17 | 22 | [17 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | [17 empty] | 2
|  |  | 0.5g | 40g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.6 | 0.55 | 0.47 | 0.49 | [17 empty] | 1
| 1 | 2 | 0.7 | 0.54 | 0.5 | 0.45 | [17 empty] | 21
| 1 | 3 | 0.6 | 0.54 | 0.5 | 0.5 | [17 empty] | 16
| 1 | 4 | 0.6 | 0.55 | 0.51 | 0.47 | [17 empty] | 0.54
| 1 | 5 | 0.6 | 0.55 | 0.55 | 0.47 | [17 empty] | 27
| 1 | 6 | 0.7 | 0.52 | 0.52 | 0.48
| Mean |  | 0.63333 | 0.54167 | 0.50833 | 0.47667
| Median |  | 0.6 | 0.545 | 0.505 | 0.475
| Std Dev. |  | 0.05164 | 0.01169 | 0.026394 | 0.017512
| Rel.Std.Dev. |  | 8.154% | 2.158% | 5.192% | 3.674%
| PDM3 |  | 17.28% | 0.3086% | -5.864% | -11.73%
| Z-Score (Absolute) |  | 2.25 | 0.35 | 0.35 | 1
| Analytical results for Eu in OREAS 926b (Indicative Value 0.6 ppm) | [22 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 08 | 15 | 17 | 22 | [17 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | [17 empty] | 2
|  |  | 0.5g | 40g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.8 | 0.584 | 0.6 | 0.44 | [17 empty] | 1
| 1 | 2 | 0.8 | 0.586 | 0.63 | 0.42 | [17 empty] | 22
| 1 | 3 | 0.7 | 0.572 | 0.61 | 0.49 | [17 empty] | 16
| 1 | 4 | 0.7 | 0.586 | 0.65 | 0.44 | [17 empty] | 0.60458
| 1 | 5 | 0.7 | 0.59 | 0.67 | 0.46 | [17 empty] | 28
| 1 | 6 | 0.8 | 0.562 | 0.67 | 0.45
| Mean |  | 0.75 | 0.58 | 0.63833 | 0.45
| Median |  | 0.75 | 0.585 | 0.64 | 0.445
| Std Dev. |  | 0.054772 | 0.010733 | 0.029944 | 0.023664
| Rel.Std.Dev. |  | 7.303% | 1.851% | 4.691% | 5.259%
| PDM3 |  | 24.05% | -4.066% | 5.582% | -25.57%
| Z-Score (Absolute) |  | 1.12 | 0.23 | 0.23 | 1.26
| Analytical results for Fe in OREAS 926b (Certified Value 3.48 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  | 3
| 1 | 1 | 3.36 | 3.43 | 3.45 | 3.39 | 3.38 | 3.18 | 3.75 | 3.59 | 3.4047 | 3.45 | 3.4027 | 3.42 | 3.43 | 3.52 | 3.17 | 3.5701 | 3.53 | 3.5394 | 3.413 |  |  | 1
| 1 | 2 | 3.44 | 3.44 | 3.48 | 3.37 | 3.39 | 3.3 | 3.62 | 3.63 | 3.4196 | 3.56 | 3.4041 | 3.47 | 3.48 | 3.57 | 3.66 | 3.4868 | 3.87 | 3.5402 | 3.406 |  |  | #N/A
| 1 | 3 | 3.46 | 3.43 | 3.48 | 3.43 | 3.38 | 2.97 | 3.61 | 3.55 | 3.4735 | 3.43 | 3.4098 | 3.42 | 3.47 | 3.51 | 3.06 | 3.6121 | 3.96 | 3.52 | 3.371 |  |  | 16
| 1 | 4 | 3.42 | 3.46 | 3.48 | 3.56 | 3.36 | 3.05 | 3.65 | 3.63 | 3.4844 | 3.46 | 3.4114 | 3.44 | 3.47 | 3.58 | 3.5 | 3.6018 | 3.92 | 3.5297 | 3.378 |  |  | 3.4751
| 1 | 5 | 3.39 | 3.46 | 3.45 | 3.43 | 3.36 | 2.97 | 3.66 | 3.53 | 3.4427 | 3.56 | 3.4103 | 3.39 | 3.48 | 3.55 | 3.61 | 3.5365 | 3.81 | 3.5684 | 3.427 |  |  | 148
| 1 | 6 | 3.5 | 3.47 | 3.38 | 3.46 | 3.35 | 3.25 | 3.53 | 3.55 | 3.4611 | 3.47 | 3.4063 | 3.43 | 3.48 | 3.56 | 3.68 | 3.5337 | 4.1 | 3.4884 | 3.364
| Mean |  | 3.4283 | 3.4483 | 3.4533 | 3.44 | 3.37 | 3.12 | 3.6367 | 3.58 | 3.4477 | 3.4883 | 3.4074 | 3.4283 | 3.4683 | 3.5483 | 3.4467 | 3.5568 | 3.865 | 3.531 | 3.3932
| Median |  | 3.43 | 3.45 | 3.465 | 3.43 | 3.37 | 3.115 | 3.635 | 3.57 | 3.4519 | 3.465 | 3.4081 | 3.425 | 3.475 | 3.555 | 3.555 | 3.5533 | 3.895 | 3.5345 | 3.392
| Std Dev. |  | 0.049967 | 0.017224 | 0.038816 | 0.066933 | 0.015492 | 0.14339 | 0.072019 | 0.043359 | 0.031163 | 0.057067 | 0.0035773 | 0.026394 | 0.019408 | 0.027869 | 0.26666 | 0.047151 | 0.19108 | 0.026425 | 0.025592
| Rel.Std.Dev. |  | 1.457% | 0.4995% | 1.124% | 1.946% | 0.4597% | 4.596% | 1.98% | 1.211% | 0.9039% | 1.636% | 0.105% | 0.7699% | 0.5596% | 0.7854% | 7.737% | 1.326% | 4.944% | 0.7484% | 0.7542%
| PDM3 |  | -1.347% | -0.7715% | -0.6276% | -1.011% | -3.026% | -10.22% | 4.648% | 3.017% | -0.7906% | 0.3796% | -1.948% | -1.347% | -0.1959% | 2.106% | -0.8194% | 2.351% | 11.22% | 1.608% | -2.359%
| Z-Score (Absolute) |  | 0.33 | 0 | 0.08 | 0.14 | 1.29 | 5.41 | 3.11 | 2.17 | 0.01 | 0.66 | 0.67 | 0.33 | 0.33 | 1.65 | 0.03 | 1.79 | 6.87 | 1.36 | 0.9
| Analytical results for Ga in OREAS 926b (Certified Value 1.67 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 1.465 | 1.52 | 1.5 | 1.51 | 1.55 | 1.81 | 6.53 | 1.9 | 1.9 | 3.2298 | 2 | < 5 | 1.82 | 0.7 | < 10 | 1.71 | 2.2812 | 1.14 |  |  |  | 1
| 1 | 2 | 1.505 | 1.49 | 1.56 | 1.46 | 1.58 | 1.98 | 7.22 | 1.9 | 1.9 | 3.3547 | 2 | < 5 | 1.79 | 0.7 | < 10 | 1.7 | 2.2084 | 1.07 |  |  |  | #N/A
| 1 | 3 | 1.52 | 1.45 | 1.46 | 1.53 | 1.6 | 1.73 | 7.49 | 2 | 2 | 2.8217 | 1.8 | < 5 | 1.84 | 0.6 | < 10 | 1.74 | 2.2159 | 1.18 |  |  |  | 16
| 1 | 4 | 1.475 | 1.46 | 1.5 | 1.53 | 1.56 | 1.79 | 6.23 | 1.8 | 1.9 | 3.025 | 2 | < 5 | 1.53 | 0.7 | < 10 | 1.71 | 2.2576 | 1.1 |  |  |  | 1.6734
| 1 | 5 | 1.5 | 1.47 | 1.46 | 1.57 | 1.42 | 1.79 | 8.69 | 1.7 | 1.9 | 3.0384 | 2 | < 5 | 1.53 | 0.6 | < 10 | 1.62 | 2.226 | 1.17 |  |  |  | 149
| 1 | 6 | 1.51 | 1.46 | 1.41 | 1.59 | 1.4 | 1.8 | 7.89 | 1.9 | 1.9 | 3.0453 | 1.8 | < 5 | 1.56 | 0.7 | < 10 | 1.62 | 2.2136 | 1.08
| Mean |  | 1.4958 | 1.475 | 1.4817 | 1.5317 | 1.5183 | 1.8167 | 7.3417 | 1.8667 | 1.9167 | 3.0858 | 1.9333 |  | 1.6783 | 0.66667 |  | 1.6833 | 2.2338 | 1.1233
| Median |  | 1.5025 | 1.465 | 1.48 | 1.53 | 1.555 | 1.795 | 7.355 | 1.9 | 1.9 | 3.0419 | 2 |  | 1.675 | 0.7 |  | 1.705 | 2.2209 | 1.12
| Std Dev. |  | 0.021311 | 0.025884 | 0.050761 | 0.045789 | 0.085888 | 0.084774 | 0.90001 | 0.10328 | 0.040825 | 0.1846 | 0.10328 |  | 0.15276 | 0.05164 |  | 0.050859 | 0.029152 | 0.046762
| Rel.Std.Dev. |  | 1.425% | 1.755% | 3.426% | 2.99% | 5.657% | 4.666% | 12.26% | 5.533% | 2.13% | 5.982% | 5.342% |  | 9.102% | 7.746% |  | 3.021% | 1.305% | 4.163%
| PDM3 |  | -10.61% | -11.86% | -11.46% | -8.472% | -9.268% | 8.559% | 338.7% | 11.55% | 14.54% | 84.4% | 15.53% |  | 0.2929% | -60.16% |  | 0.5917% | 33.48% | -32.87%
| Z-Score (Absolute) |  | 0.66 | 0.71 | 0.69 | 0.56 | 0.6 | 0.17 | 14.43 | 0.3 | 0.43 | 3.45 | 0.47 | 1.94 | 0.18 | 2.8 | 8.39 | 0.17 | 1.25 | 1.62
| Analytical results for Gd in OREAS 926b (Certified Value 2.81 ppm) | [22 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 14 | 15 | 17 | 22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [16 empty] | 2
|  |  | 0.5g | 0.25g | 40g | 0.5g | 0.5g | [16 empty] | 2
| 1 | 1 | 3.3 | 3.6199 | 2.74 | 2.36 | 2.13 | [16 empty] | 1
| 1 | 2 | 3.4 | 3.6719 | 2.73 | 2.33 | 2.03 | [16 empty] | 6
| 1 | 3 | 3.1 | 3.5646 | 2.68 | 2.35 | 2.28 | [16 empty] | 16
| 1 | 4 | 3.3 | 3.5722 | 2.76 | 2.32 | 2.15 | [16 empty] | 2.8116
| 1 | 5 | 3.2 | 3.4741 | 2.78 | 2.32 | 2.22 | [16 empty] | 150
| 1 | 6 | 3.4 | 3.6054 | 2.58 | 2.23 | 2.15
| Mean |  | 3.2833 | 3.5847 | 2.7117 | 2.3183 | 2.16
| Median |  | 3.3 | 3.5888 | 2.735 | 2.325 | 2.15
| Std Dev. |  | 0.1169 | 0.06644 | 0.072778 | 0.046224 | 0.084853
| Rel.Std.Dev. |  | 3.561% | 1.853% | 2.684% | 1.994% | 3.928%
| PDM3 |  | 16.78% | 27.5% | -3.554% | -17.54% | -23.18%
| Z-Score (Absolute) |  | 0.7 | 1.07 | 0 | 0.48 | 0.67
| Analytical results for Ge in OREAS 926b (Certified Value 0.1 ppm) | [22 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 17 | 18 | 20 | 21 | 22 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.059 | 0.07 | 0.08 | 0.11 | 0.12 | < 0.1 | 0.12 | 0.21 | < 0.1 | 0.09 | 0.14188 | < 0.01 | [9 empty] | 1
| 1 | 2 | 0.063 | 0.08 | 0.08 | 0.11 | 0.12 | < 0.1 | 0.13 | 0.2 | < 0.1 | 0.1 | 0.13363 | < 0.01 | [9 empty] | 17
| 1 | 3 | 0.063 | 0.07 | 0.1 | 0.11 | 0.11 | < 0.1 | 0.12 | 0.19 | < 0.1 | 0.11 | 0.15238 | < 0.01 | [9 empty] | 16
| 1 | 4 | 0.061 | 0.07 | 0.11 | 0.11 | 0.09 | < 0.1 | 0.12 | 0.18 | < 0.1 | 0.1 | 0.14508 | < 0.01 | [9 empty] | 0.09959
| 1 | 5 | 0.061 | 0.07 | 0.1 | 0.1 | 0.08 | < 0.1 | 0.1 | 0.18 | < 0.1 | 0.1 | 0.16127 | < 0.01 | [9 empty] | 151
| 1 | 6 | 0.057 | 0.08 | 0.1 | 0.11 | 0.07 | < 0.1 | 0.11 | 0.18 | < 0.1 | 0.09 | 0.14208 | < 0.01
| Mean |  | 0.060667 | 0.073333 | 0.095 | 0.10833 | 0.098333 |  | 0.11667 | 0.19 |  | 0.098333 | 0.14605
| Median |  | 0.061 | 0.07 | 0.1 | 0.11 | 0.1 |  | 0.12 | 0.185 |  | 0.1 | 0.14358
| Std Dev. |  | 0.0023381 | 0.005164 | 0.012247 | 0.0040825 | 0.02137 |  | 0.010328 | 0.012649 |  | 0.0075277 | 0.0095938
| Rel.Std.Dev. |  | 3.854% | 7.042% | 12.89% | 3.768% | 21.73% |  | 8.853% | 6.657% |  | 7.655% | 6.569%
| PDM3 |  | -39.08% | -26.36% | -4.609% | 8.78% | -1.262% |  | 17.15% | 90.78% |  | -1.262% | 46.65%
| Z-Score (Absolute) |  | 0.82 | 0.53 | 0.04 | 0.27 | 0.04 | 1.06 | 0.45 | 2.12 | 1.06 | 0.04 | 1.12 | 2.08
| Analytical results for Hf in OREAS 926b (Certified Value 0.29 ppm) | [22 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 17 | 18 | 20 | 22 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.269 | 0.3 | 0.27 | 0.27 | 0.27 | 0.2 | 0.29 | 0.28 | 0.45 | 0.35 | 0.27 | 0.26 | [9 empty] | 1
| 1 | 2 | 0.289 | 0.31 | 0.29 | 0.27 | 0.28 | 0.2 | 0.29 | 0.29 | 0.47 | 0.3 | 0.27 | 0.23 | [9 empty] | 13
| 1 | 3 | 0.259 | 0.31 | 0.26 | 0.27 | 0.27 | 0.2 | 0.29 | 0.29 | 0.48 | 0.28 | 0.31 | 0.23 | [9 empty] | 16
| 1 | 4 | 0.278 | 0.31 | 0.27 | 0.28 | 0.28 | 0.2 | 0.29 | 0.29 | 0.44 | 0.33 | 0.31 | 0.23 | [9 empty] | 0.28674
| 1 | 5 | 0.269 | 0.3 | 0.26 | 0.29 | 0.29 | 0.2 | 0.3 | 0.28 | 0.43 | 0.29 | 0.3 | 0.24 | [9 empty] | 152
| 1 | 6 | 0.282 | 0.3 | 0.27 | 0.28 | 0.29 | 0.2 | 0.27 | 0.27 | 0.42 | 0.34 | 0.31 | 0.23
| Mean |  | 0.27433 | 0.305 | 0.27 | 0.27667 | 0.28 | 0.2 | 0.28833 | 0.28333 | 0.44833 | 0.315 | 0.295 | 0.23667
| Median |  | 0.2735 | 0.305 | 0.27 | 0.275 | 0.28 | 0.2 | 0.29 | 0.285 | 0.445 | 0.315 | 0.305 | 0.23
| Std Dev. |  | 0.010764 | 0.0054772 | 0.010954 | 0.008165 | 0.0089443 | 3.0405e-17 | 0.0098319 | 0.008165 | 0.023166 | 0.02881 | 0.019748 | 0.012111
| Rel.Std.Dev. |  | 3.924% | 1.796% | 4.057% | 2.951% | 3.194% | 1.52e-14% | 3.41% | 2.882% | 5.167% | 9.146% | 6.694% | 5.117%
| PDM3 |  | -4.327% | 6.368% | -5.838% | -3.513% | -2.351% | -30.25% | 0.5554% | -1.188% | 56.35% | 9.855% | 2.88% | -17.46%
| Z-Score (Absolute) |  | 0.52 | 1.25 | 0.77 | 0.39 | 0.19 | NA | 0.29 | 0 | 9.54 | 1.83 | 0.67 | 2.7
Results from laboratory 8 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 926b (Certified Value 0.039 ppm) | [22 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 20 | 22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.048 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.11 | < 0.005 | 0.27 | < 0.5 | < 0.05 | 0.29 | 0.02 | 0.13 | 0.14 | [6 empty] | 1
| 1 | 2 | 0.043 | 0.05 | 0.04 | 0.02 | 0.04 | 0.04 | 0.05 | < 0.005 | 0.13 | < 0.5 | < 0.05 | 0.27 | < 0.01 | 0.1 | 0.12 | [6 empty] | 16
| 1 | 3 | 0.036 | 0.04 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 | 0.022 | 0.12 | < 0.5 | < 0.05 | 0.28 | < 0.01 | 0.07 | 0.14 | [6 empty] | 16
| 1 | 4 | 0.039 | 0.05 | 0.04 | 0.03 | 0.03 | 0.05 | 0.03 | < 0.005 | 0.12 | < 0.5 | 0.05 | 0.29 | < 0.01 | 0.08 | 0.1 | [6 empty] | 0.038648
| 1 | 5 | 0.035 | 0.06 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.005 | 0.11 | < 0.5 | < 0.05 | 0.29 | < 0.01 | 0.07 | 0.11 | [6 empty] | 153
| 1 | 6 | 0.036 | 0.05 | 0.04 | 0.03 | 0.03 | 0.04 | 0.04 | < 0.005 | 0.1 | < 0.5 | < 0.05 | 0.29 | < 0.01 | 0.07 | 0.08
| Mean |  | 0.0395 | 0.048333 | 0.04 | 0.033333 | 0.035 | 0.041667 | 0.051667 | 0.0135 | 0.14167 |  | 0.05 | 0.285 | 0.02 | 0.086667 | 0.115
| Median |  | 0.0375 | 0.05 | 0.04 | 0.035 | 0.035 | 0.04 | 0.04 | 0.0135 | 0.12 |  | 0.05 | 0.29 | 0.02 | 0.075 | 0.115
| Std Dev. |  | 0.0050892 | 0.0075277 | 0 | 0.008165 | 0.0054772 | 0.0075277 | 0.029269 | 0.012021 | 0.063692 |  |  | 0.0083666 |  | 0.024221 | 0.023452
| Rel.Std.Dev. |  | 12.88% | 15.57% | 0% | 24.49% | 15.65% | 18.07% | 56.65% | 89.04% | 44.96% |  |  | 2.936% |  | 27.95% | 20.39%
| PDM3 |  | 2.204% | 25.06% | 3.498% | -13.75% | -9.439% | 7.81% | 33.68% | -65.07% | 266.6% |  | 29.37% | 637.4% | -48.25% | 124.2% | 197.6%
| Z-Score (Absolute) |  | 0.12 | 0.36 | 0.09 | 0.45 | 0.36 | 0 | 0.54 | 1.92 | 5.39 | 11.24 | 0.67 | 13.13 | 1.84 | 2.43 | 3.96
| Analytical results for Ho in OREAS 926b (Indicative Value 0.23 ppm) | [22 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 08 | 15 | 17 | 22 | [17 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | [17 empty] | 2
|  |  | 0.5g | 40g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.3 | 0.255 | 0.21 | 0.18 | [17 empty] | 1
| 1 | 2 | 0.3 | 0.265 | 0.22 | 0.17 | [17 empty] | 23
| 1 | 3 | 0.2 | 0.25 | 0.21 | 0.19 | [17 empty] | 16
| 1 | 4 | 0.3 | 0.26 | 0.2 | 0.19 | [17 empty] | 0.23312
| 1 | 5 | 0.3 | 0.265 | 0.2 | 0.19 | [17 empty] | 29
| 1 | 6 | 0.3 | 0.25 | 0.2 | 0.19
| Mean |  | 0.28333 | 0.2575 | 0.20667 | 0.185
| Median |  | 0.3 | 0.2575 | 0.205 | 0.19
| Std Dev. |  | 0.040825 | 0.006892 | 0.008165 | 0.0083666
| Rel.Std.Dev. |  | 14.41% | 2.677% | 3.951% | 4.522%
| PDM3 |  | 21.54% | 10.46% | -11.35% | -20.64%
| Z-Score (Absolute) |  | 0.95 | 0.47 | 0.47 | 0.88
| Analytical results for In in OREAS 926b (Certified Value 0.72 ppm) | [22 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 15 | 17 | 18 | 20 | 21 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | [8 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 0.5g | 0.25g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | 0.676 | 0.749 | 0.724 | 0.697 | 0.728 | 0.73 | 0.747 | 0.71 | 1 | 0.71 | 0.74 | 0.667 | 0.69039 | [8 empty] | 1
| 1 | 2 | 0.713 | 0.756 | 0.714 | 0.674 | 0.725 | 0.75 | 0.774 | 0.71 | 1 | 0.73 | 0.85 | 0.673 | 0.68031 | [8 empty] | #N/A
| 1 | 3 | 0.69 | 0.756 | 0.76 | 0.715 | 0.732 | 0.7 | 0.736 | 0.7 | 1 | 0.713 | 0.74 | 0.676 | 0.69393 | [8 empty] | 16
| 1 | 4 | 0.725 | 0.759 | 0.744 | 0.726 | 0.736 | 0.77 | 0.755 | 0.7 | 1 | 0.664 | 0.82 | 0.672 | 0.69594 | [8 empty] | 0.72106
| 1 | 5 | 0.696 | 0.754 | 0.722 | 0.734 | 0.686 | 0.75 | 0.741 | 0.71 | 1 | 0.638 | 0.83 | 0.653 | 0.69543 | [8 empty] | 154
| 1 | 6 | 0.716 | 0.751 | 0.718 | 0.739 | 0.7 | 0.81 | 0.717 | 0.7 | 1 | 0.649 | 0.82 | 0.661 | 0.68934
| Mean |  | 0.70267 | 0.75417 | 0.73033 | 0.71417 | 0.71783 | 0.75167 | 0.745 | 0.705 | 1 | 0.684 | 0.8 | 0.667 | 0.69089
| Median |  | 0.7045 | 0.755 | 0.723 | 0.7205 | 0.7265 | 0.75 | 0.744 | 0.705 | 1 | 0.687 | 0.82 | 0.6695 | 0.69216
| Std Dev. |  | 0.018435 | 0.003656 | 0.017862 | 0.024734 | 0.020084 | 0.037103 | 0.01911 | 0.0054772 | 0 | 0.038403 | 0.047749 | 0.0086487 | 0.0058311
| Rel.Std.Dev. |  | 2.624% | 0.4848% | 2.446% | 3.463% | 2.798% | 4.936% | 2.565% | 0.7769% | 0% | 5.614% | 5.969% | 1.297% | 0.844%
| PDM3 |  | -2.551% | 4.591% | 1.286% | -0.9561% | -0.4475% | 4.245% | 3.32% | -2.227% | 38.68% | -5.14% | 10.95% | -7.497% | -4.184%
| Z-Score (Absolute) |  | 0.33 | 0.95 | 0.36 | 0.05 | 0.05 | 0.89 | 0.72 | 0.27 | NA | 0.8 | 2.09 | 1.22 | 0.63
Results from laboratory 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 926b (Certified Value 0.174 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  |  | 3
| 1 | 1 | 0.17 | 0.17 | 0.19 | 0.17 | 0.17 | 0.23 | 0.17 | 0.18 | 0.225 | 0.22664 | 0.19 | 0.21 | 0.17 | 0.12 | 0.21269 | 0.15 | 0.259 | 0.133 |  |  |  | 1
| 1 | 2 | 0.17 | 0.16 | 0.19 | 0.16 | 0.17 | 0.21 | 0.17 | 0.18 | 0.23 | 0.22255 | 0.19 | 0.21 | 0.17 | 0.14 | 0.20866 | 0.16 | 0.261 | 0.125 |  |  |  | #N/A
| 1 | 3 | 0.17 | 0.16 | 0.17 | 0.17 | 0.17 | 0.22 | 0.17 | 0.17 | 0.216 | 0.22584 | 0.18 | 0.21 | 0.18 | 0.12 | 0.2096 | 0.16 | 0.259 | 0.133 |  |  |  | 16
| 1 | 4 | 0.17 | 0.16 | 0.17 | 0.17 | 0.17 | 0.23 | 0.17 | 0.16 | 0.225 | 0.22301 | 0.18 | 0.22 | 0.17 | 0.14 | 0.21516 | 0.17 | 0.262 | 0.125 |  |  |  | 0.17412
| 1 | 5 | 0.17 | 0.16 | 0.17 | 0.17 | 0.16 | 0.23 | 0.17 | 0.16 | 0.239 | 0.22711 | 0.19 | 0.22 | 0.17 | 0.14 | 0.21265 | 0.16 | 0.264 | 0.125 |  |  |  | 155
| 1 | 6 | 0.17 | 0.17 | 0.16 | 0.17 | 0.16 | 0.26 | 0.16 | 0.18 | 0.241 | 0.22765 | 0.18 | 0.21 | 0.18 | 0.14 | 0.21247 | 0.18 | 0.2541 | 0.125
| Mean |  | 0.17 | 0.16333 | 0.175 | 0.16833 | 0.16667 | 0.23 | 0.16833 | 0.17167 | 0.22933 | 0.22546 | 0.185 | 0.21333 | 0.17333 | 0.13333 | 0.21187 | 0.16333 | 0.25985 | 0.12767
| Median |  | 0.17 | 0.16 | 0.17 | 0.17 | 0.17 | 0.23 | 0.17 | 0.175 | 0.2275 | 0.22624 | 0.185 | 0.21 | 0.17 | 0.14 | 0.21256 | 0.16 | 0.26 | 0.125
| Std Dev. |  | 0 | 0.005164 | 0.012247 | 0.0040825 | 0.005164 | 0.016733 | 0.0040825 | 0.0098319 | 0.0094375 | 0.0021681 | 0.0054772 | 0.005164 | 0.005164 | 0.010328 | 0.002364 | 0.010328 | 0.0033963 | 0.0041312
| Rel.Std.Dev. |  | 0% | 3.162% | 6.999% | 2.425% | 3.098% | 7.275% | 2.425% | 5.727% | 4.115% | 0.9616% | 2.961% | 2.421% | 2.979% | 7.746% | 1.116% | 6.323% | 1.307% | 3.236%
| PDM3 |  | -2.365% | -6.194% | 0.5064% | -3.322% | -4.28% | 32.09% | -3.322% | -1.408% | 31.71% | 29.49% | 6.25% | 22.52% | -0.4508% | -23.42% | 21.68% | -6.194% | 49.24% | -26.68%
| Z-Score (Absolute) |  | 0.16 | 0.57 | 0.16 | 0.26 | 0.36 | 3.58 | 0.26 | 0.05 | 3.54 | 3.3 | 0.78 | 2.54 | 0.05 | 2.44 | 2.45 | 0.57 | 5.44 | 2.81
| Analytical results for La in OREAS 926b (Certified Value 22.4 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 21.7 | 22.7 | 24.5 | 22.6 | 22.8 | 29.4 | 16.7 | 19.9 | 26.3 | 23.9906 | 31.8 | 27.8 | 22.83 | 12 | 25.73 | 17.2 | 13 |  |  |  |  | 1
| 1 | 2 | 22.7 | 22 | 24.9 | 21.7 | 23.9 | 30.5 | 16.7 | 20 | 26.9 | 23.767 | 31.4 | 29 | 23.15 | 14.6 | 24.578 | 17.4 | 14 |  |  |  |  | #N/A
| 1 | 3 | 21.8 | 22 | 23 | 21.4 | 22.8 | 27 | 16.5 | 19 | 26.4 | 24.0331 | 31.4 | 28 | 23.2 | 13.8 | 26.058 | 17.7 | 14 |  |  |  |  | 16
| 1 | 4 | 22.3 | 22.4 | 24 | 21.5 | 23 | 29.1 | 16.5 | 17.7 | 26.5 | 23.3726 | 31.6 | 28.6 | 20.99 | 15.2 | 27.246 | 18.2 | 14 |  |  |  |  | 22.4408
| 1 | 5 | 22.3 | 21.4 | 22.8 | 22.7 | 22.5 | 28.1 | 16.7 | 18.2 | 27 | 23.6686 | 31.6 | 28.6 | 20.62 | 15.3 | 25.518 | 17.6 | 14 |  |  |  |  | 156
| 1 | 6 | 23.5 | 22.6 | 22.7 | 22.3 | 22.6 | 30.8 | 15.9 | 19.7 | 26.5 | 23.3415 | 30.2 | 27.8 | 20.31 | 15.2 | 25.641 | 17 | 13
| Mean |  | 22.3833 | 22.1833 | 23.65 | 22.0333 | 22.9333 | 29.15 | 16.5 | 19.0833 | 26.6 | 23.6956 | 31.3333 | 28.3 | 21.85 | 14.35 | 25.7952 | 17.5167 | 13.6667
| Median |  | 22.3 | 22.2 | 23.5 | 22 | 22.8 | 29.25 | 16.6 | 19.35 | 26.5 | 23.7178 | 31.5 | 28.3 | 21.91 | 14.9 | 25.6855 | 17.5 | 14
| Std Dev. |  | 0.65853 | 0.48339 | 0.94393 | 0.57155 | 0.50465 | 1.4377 | 0.30984 | 0.95795 | 0.28284 | 0.29549 | 0.57504 | 0.502 | 1.3489 | 1.2834 | 0.86722 | 0.42151 | 0.5164
| Rel.Std.Dev. |  | 2.942% | 2.179% | 3.991% | 2.594% | 2.2% | 4.932% | 1.878% | 5.02% | 1.063% | 1.247% | 1.835% | 1.774% | 6.173% | 8.943% | 3.362% | 2.406% | 3.779%
| PDM3 |  | -0.2562% | -1.147% | 5.388% | -1.816% | 2.195% | 29.9% | -26.47% | -14.96% | 18.53% | 5.591% | 39.63% | 26.11% | -2.633% | -36.05% | 14.95% | -21.94% | -39.1%
| Z-Score (Absolute) |  | 0 | 0.04 | 0.25 | 0.07 | 0.11 | 1.34 | 1.16 | 0.65 | 0.83 | 0.26 | 1.77 | 1.17 | 0.11 | 1.59 | 0.67 | 0.96 | 1.72
| Analytical results for Li in OREAS 926b (Certified Value 1.61 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 1.7 | 1 | 1.8 | 1.6 | 1.6 | 2.1 | 1.6 | 2.2 | 1.9936 | < 10 | 1.41 | < 1 | < 10 | 1.2 | < 1 | [6 empty] | 1
| 1 | 2 | 1.8 | 1.1 | 1.9 | 1.6 | 1.7 | 2.2 | 1.7 | 2.2 | 2.3214 | < 10 | 1.42 | < 1 | < 10 | 1.1 | < 1 | [6 empty] | #N/A
| 1 | 3 | 1.7 | 1.1 | 1.5 | 1.6 | 1.7 | 2.1 | 1.6 | 2.2 | 1.978 | < 10 | 1.58 | < 1 | < 10 | 1.2 | < 1 | [6 empty] | 16
| 1 | 4 | 1.7 | 1 | 1.6 | 1.6 | 1.7 | 2.1 | 1.6 | 2.2 | 2.1068 | < 10 | 1.26 | < 1 | < 10 | 1.1 | 1 | [6 empty] | 1.6108
| 1 | 5 | 1.7 | 1.1 | 1.5 | 1.6 | 1.5 | 2.1 | 1.8 | 2.4 | 2.0458 | < 10 | 1.29 | < 1 | < 10 | 1.1 | < 1 | [6 empty] | 157
| 1 | 6 | 1.8 | 1.1 | 1.4 | 1.6 | 1.6 | 2.4 | 1.7 | 2.4 | 1.923 | < 10 | 1.35 | < 1 | < 10 | 1.1 | < 1
| Mean |  | 1.7333 | 1.0667 | 1.6167 | 1.6 | 1.6333 | 2.1667 | 1.6667 | 2.2667 | 2.0614 |  | 1.385 |  |  | 1.1333 | 1
| Median |  | 1.7 | 1.1 | 1.55 | 1.6 | 1.65 | 2.1 | 1.65 | 2.2 | 2.0197 |  | 1.38 |  |  | 1.1 | 1
| Std Dev. |  | 0.05164 | 0.05164 | 0.19408 | 2.4324e-16 | 0.08165 | 0.12111 | 0.08165 | 0.10328 | 0.14184 |  | 0.11467 |  |  | 0.05164
| Rel.Std.Dev. |  | 2.979% | 4.841% | 12% | 1.52e-14% | 4.999% | 5.59% | 4.899% | 4.556% | 6.881% |  | 8.28% |  |  | 4.556%
| PDM3 |  | 7.606% | -33.78% | 0.3634% | -0.6713% | 1.398% | 34.51% | 3.467% | 40.72% | 27.97% |  | -14.02% |  |  | -29.64% | -37.92%
| Z-Score (Absolute) |  | 0.13 | 0.76 | 0.02 | 0.04 | 0 | 0.72 | 0.04 | 0.85 | 0.58 | 4.54 | 0.33 | 1.53 | 4.54 | 0.67 | 1.42
| Analytical results for Lu in OREAS 926b (Indicative Value 0.068 ppm) | [22 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 08 | 15 | 17 | 18 | [17 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [17 empty] | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | [17 empty] | 3
| 1 | 1 | < 0.1 | 0.08 | 0.06 | 0.05 | [17 empty] | 1
| 1 | 2 | < 0.1 | 0.08 | 0.07 | 0.06 | [17 empty] | 24
| 1 | 3 | < 0.1 | 0.08 | 0.07 | 0.05 | [17 empty] | 16
| 1 | 4 | < 0.1 | 0.08 | 0.06 | 0.06 | [17 empty] | 0.0675
| 1 | 5 | < 0.1 | 0.08 | 0.07 | 0.06 | [17 empty] | 30
| 1 | 6 | < 0.1 | 0.075 | 0.07 | 0.06
| Mean |  |  | 0.079167 | 0.066667 | 0.056667
| Median |  |  | 0.08 | 0.07 | 0.06
| Std Dev. |  |  | 0.0020412 | 0.005164 | 0.005164
| Rel.Std.Dev. |  |  | 2.578% | 7.746% | 9.113%
| PDM3 |  |  | 17.28% | -1.235% | -16.05%
| Z-Score (Absolute) |  | 0.94 | 1.42 | 0.4 | 0.4
| Analytical results for Mg in OREAS 926b (Certified Value 0.047 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  |  | 3
| 1 | 1 | 0.04 | 0.04 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.063666 | 0.05 | 0.05 | 0.05 | 0.03 | 0.053703 | 0.04 | 0.054 | 0.042 |  |  |  | 1
| 1 | 2 | 0.04 | 0.04 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.063683 | 0.05 | 0.05 | 0.05 | 0.04 | 0.055579 | 0.04 | 0.054 | 0.042 |  |  |  | #N/A
| 1 | 3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.05 | 0.063692 | 0.04 | 0.05 | 0.05 | 0.03 | 0.053764 | 0.04 | 0.054 | 0.042 |  |  |  | 16
| 1 | 4 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.05 | 0.063615 | 0.05 | 0.05 | 0.06 | 0.04 | 0.054297 | 0.04 | 0.054 | 0.042 |  |  |  | 0.046893
| 1 | 5 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.05 | 0.063757 | 0.05 | 0.05 | 0.06 | 0.04 | 0.053529 | 0.04 | 0.0594 | 0.042 |  |  |  | 158
| 1 | 6 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.06 | 0.04 | 0.05 | 0.05 | 0.063788 | 0.05 | 0.05 | 0.05 | 0.04 | 0.054621 | 0.04 | 0.0581 | 0.042
| Mean |  | 0.041667 | 0.04 | 0.043333 | 0.04 | 0.04 | 0.051667 | 0.048333 | 0.045 | 0.05 | 0.0637 | 0.048333 | 0.05 | 0.053333 | 0.036667 | 0.054249 | 0.04 | 0.055583 | 0.042
| Median |  | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | 0.045 | 0.05 | 0.063687 | 0.05 | 0.05 | 0.05 | 0.04 | 0.054031 | 0.04 | 0.054 | 0.042
| Std Dev. |  | 0.0040825 | 0 | 0.005164 | 0 | 0 | 0.0040825 | 0.0040825 | 0.0054772 | 7.6012e-18 | 6.2699e-05 | 0.0040825 | 7.6012e-18 | 0.005164 | 0.005164 | 0.00076965 | 0 | 0.0024871 | 0
| Rel.Std.Dev. |  | 9.798% | 0% | 11.92% | 0% | 0% | 7.902% | 8.447% | 12.17% | 1.52e-14% | 0.09843% | 8.447% | 1.52e-14% | 9.682% | 14.08% | 1.419% | 0% | 4.475% | 0%
| PDM3 |  | -11.15% | -14.7% | -7.592% | -14.7% | -14.7% | 10.18% | 3.071% | -4.038% | 6.625% | 35.84% | 3.071% | 6.625% | 13.73% | -21.81% | 15.69% | -14.7% | 18.53% | -10.44%
| Z-Score (Absolute) |  | 0.58 | 0.77 | 0.39 | 0.77 | 0.77 | 0.58 | 0.19 | 0.19 | 0.39 | 1.97 | 0.19 | 0.39 | 0.77 | 1.16 | 0.88 | 0.77 | 1.03 | 0.51
| Analytical results for Mn in OREAS 926b (Certified Value 0.005 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.00514 | 0.0052 | 0.0052 | 0.0051 | 0.0049 | 0.0046 | 0.0056 | 0.0055 | 0.0054 | 0.006 | 0.0053 | 0.0055 | 0.004 | 0.0051234 | 0.0049 | 0.0050183 | 0.0054 |  |  |  |  | 1
| 1 | 2 | 0.00525 | 0.0052 | 0.0053 | 0.0051 | 0.005 | 0.0051 | 0.0045 | 0.0056 | 0.0053 | 0.005 | 0.0053 | 0.0054 | 0.0046 | 0.0049442 | 0.0051 | 0.005009 | 0.0054 |  |  |  |  | #N/A
| 1 | 3 | 0.0052 | 0.0051 | 0.0052 | 0.0051 | 0.0052 | 0.0044 | 0.0043 | 0.0054 | 0.0052 | 0.005 | 0.0054 | 0.0055 | 0.0039 | 0.0050224 | 0.0052 | 0.0050741 | 0.0054 |  |  |  |  | 16
| 1 | 4 | 0.00512 | 0.0052 | 0.0052 | 0.0052 | 0.0049 | 0.0046 | 0.0042 | 0.0054 | 0.0053 | 0.005 | 0.0054 | 0.0052 | 0.0044 | 0.0050898 | 0.0051 | 0.0050516 | 0.0054 |  |  |  |  | 0.005155
| 1 | 5 | 0.00502 | 0.0052 | 0.0052 | 0.0051 | 0.005 | 0.0045 | 0.0041 | 0.0053 | 0.0054 | 0.005 | 0.0052 | 0.0051 | 0.0045 | 0.0049959 | 0.0051 | 0.0050386 | 0.0054 |  |  |  |  | 159
| 1 | 6 | 0.00525 | 0.0052 | 0.005 | 0.0051 | 0.0051 | 0.0049 | 0.0039 | 0.0054 | 0.0052 | 0.005 | 0.0053 | 0.0052 | 0.0047 | 0.0049774 | 0.005 | 0.0050552 | 0.0054
| Mean |  | 0.0051633 | 0.0051833 | 0.0051833 | 0.0051167 | 0.0050167 | 0.0046833 | 0.0044333 | 0.0054333 | 0.0053 | 0.0051667 | 0.0053167 | 0.0053167 | 0.00435 | 0.0050255 | 0.0050667 | 0.0050411 | 0.0054
| Median |  | 0.00517 | 0.0052 | 0.0052 | 0.0051 | 0.005 | 0.0046 | 0.00425 | 0.0054 | 0.0053 | 0.005 | 0.0053 | 0.0053 | 0.00445 | 0.0050092 | 0.0051 | 0.0050451 | 0.0054
| Std Dev. |  | 8.8694e-05 | 4.0825e-05 | 9.8319e-05 | 4.0825e-05 | 0.0001169 | 0.00026394 | 0.00060553 | 0.00010328 | 8.9443e-05 | 0.00040825 | 7.5277e-05 | 0.00017224 | 0.00032711 | 6.8597e-05 | 0.00010328 | 2.4308e-05 | 9.5015e-19
| Rel.Std.Dev. |  | 1.718% | 0.7876% | 1.897% | 0.7979% | 2.33% | 5.636% | 13.66% | 1.901% | 1.688% | 7.902% | 1.416% | 3.24% | 7.52% | 1.365% | 2.038% | 0.4822% | 1.76e-14%
| PDM3 |  | 0.1623% | 0.5503% | 0.5503% | -0.743% | -2.683% | -9.149% | -14% | 5.4% | 2.813% | 0.227% | 3.137% | 3.137% | -15.62% | -2.511% | -1.713% | -2.208% | 4.753%
| Z-Score (Absolute) |  | 0 | 0.1 | 0.1 | 0.23 | 0.72 | 2.35 | 3.57 | 1.32 | 0.67 | 0.02 | 0.75 | 0.75 | 3.98 | 0.67 | 0.47 | 0.6 | 1.26
| Analytical results for Mo in OREAS 926b (Certified Value 4.77 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 4.54 | 4.44 | 4.56 | 4.65 | 4.86 | 4.86 | 4.94 | 4.66 | 5.2 | 4.5 | 6 | 4.9 | 5.51 | 4.95 | < 5 | 4.77 | 4.8239 | 4.89 |  |  |  | 1
| 1 | 2 | 4.73 | 4.52 | 4.56 | 4.65 | 4.87 | 5.08 | 4.22 | 4.71 | 5.1 | 4.5 | 6.2 | 4.9 | 5.54 | 5.59 | < 5 | 4.82 | 4.6694 | 4.82 |  |  |  | 1
| 1 | 3 | 4.58 | 4.44 | 4.57 | 4.36 | 4.87 | 4.48 | 4.32 | 4.51 | 5.1 | 4.7 | 5.8 | 4.9 | 5.17 | 4.8 | < 5 | 4.7 | 4.8665 | 5.09 |  |  |  | 16
| 1 | 4 | 4.68 | 4.48 | 4.66 | 4.42 | 4.89 | 4.7 | 4.32 | 4.84 | 5.2 | 4.6 | 5.6 | 4.9 | 4.93 | 5.34 | < 5 | 4.75 | 4.752 | 4.75 |  |  |  | 4.7689
| 1 | 5 | 4.7 | 4.43 | 4.55 | 4.59 | 4.57 | 4.7 | 4.31 | 4.61 | 5.2 | 4.7 | 5.8 | 4.8 | 4.86 | 5.41 | < 5 | 4.55 | 4.8696 | 5.06 |  |  |  | 160
| 1 | 6 | 4.79 | 4.46 | 4.49 | 4.48 | 4.55 | 4.99 | 3.98 | 4.58 | 5.3 | 4.7 | 6 | 4.9 | 4.87 | 5.54 | < 5 | 4.65 | 4.7818 | 4.84
| Mean |  | 4.67 | 4.4617 | 4.565 | 4.525 | 4.7683 | 4.8017 | 4.3483 | 4.6517 | 5.1833 | 4.6167 | 5.9 | 4.8833 | 5.1467 | 5.2717 |  | 4.7067 | 4.7938 | 4.9083
| Median |  | 4.69 | 4.45 | 4.56 | 4.535 | 4.865 | 4.78 | 4.315 | 4.635 | 5.2 | 4.65 | 5.9 | 4.9 | 5.05 | 5.375 |  | 4.725 | 4.8028 | 4.865
| Std Dev. |  | 0.093808 | 0.033714 | 0.054681 | 0.12309 | 0.16179 | 0.21931 | 0.31789 | 0.11479 | 0.075277 | 0.098319 | 0.20976 | 0.040825 | 0.31399 | 0.32344 |  | 0.096471 | 0.07654 | 0.13703
| Rel.Std.Dev. |  | 2.009% | 0.7556% | 1.198% | 2.72% | 3.393% | 4.567% | 7.311% | 2.468% | 1.452% | 2.13% | 3.555% | 0.836% | 6.101% | 6.136% |  | 2.05% | 1.597% | 2.792%
| PDM3 |  | -2.074% | -6.442% | -4.275% | -5.114% | -0.01159% | 0.6874% | -8.819% | -2.458% | 8.691% | -3.192% | 23.72% | 2.4% | 7.922% | 10.54% |  | -1.305% | 0.5233% | 2.924%
| Z-Score (Absolute) |  | 0.27 | 1.08 | 0.68 | 0.83 | 0.12 | 0.25 | 1.53 | 0.34 | 1.75 | 0.47 | 4.57 | 0.57 | 1.61 | 2.1 | 8.79 | 0.12 | 0.22 | 0.67
| Analytical results for Na in OREAS 926b (Indicative Value 0.011 wt.%) | [22 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.016 | 0.01 | 0.006 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0099837 | 0.01 | 0.018 | < 0.007 |  |  |  |  | 1
| 1 | 2 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.016 | < 0.01 | 0.007 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0096147 | 0.01 | 0.018 | < 0.007 |  |  |  |  | 14
| 1 | 3 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.016 | < 0.01 | 0.006 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0098699 | 0.01 | 0.017 | < 0.007 |  |  |  |  | 16
| 1 | 4 | 0.011 | 0.01 | 0.01 | 0.01 | 0.01 | 0.015 | < 0.01 | 0.006 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.010037 | 0.01 | 0.018 | < 0.007 |  |  |  |  | 0.010765
| 1 | 5 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.016 | 0.01 | 0.006 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0098356 | 0.01 | 0.031 | < 0.007 |  |  |  |  | 25
| 1 | 6 | 0.011 | 0.01 | 0.01 | 0.01 | 0.01 | 0.016 | < 0.01 | 0.006 | 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.0098145 | 0.01 | 0.026 | < 0.007
| Mean |  | 0.010333 | 0.01 | 0.01 | 0.01 | 0.01 | 0.015833 | 0.01 | 0.0061667 | 0.01 |  |  | 0.01 |  | 0.0098593 | 0.01 | 0.021333
| Median |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.016 | 0.01 | 0.006 | 0.01 |  |  | 0.01 |  | 0.0098528 | 0.01 | 0.018
| Std Dev. |  | 0.0005164 | 0 | 0 | 0 | 0 | 0.00040825 | 0 | 0.00040825 | 0 |  |  | 0 |  | 0.00014813 | 0 | 0.005785
| Rel.Std.Dev. |  | 4.997% | 0% | 0% | 0% | 0% | 2.578% | 0% | 6.62% | 0% |  |  | 0% |  | 1.502% | 0% | 27.12%
| PDM3 |  | -4.008% | -7.105% | -7.105% | -7.105% | -7.105% | 47.08% | -7.105% | -42.71% | -7.105% |  |  | -7.105% |  | -8.412% | -7.105% | 98.18%
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 0 | 0 | 11.8 | 6.74 | 7.75 | 0 | 10.11 | 10.11 | 0 | 10.11 | 0.28 | 0 | 22.93 | 12.73
| Analytical results for Nb in OREAS 926b (Indicative Value 0.044 ppm) | [22 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 17 | 18 | 20 | 22 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.026 | 0.05 | 0.08 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | 0.43 | 0.85 | 0.08 | 0.03 | [9 empty] | 1
| 1 | 2 | 0.026 | 0.05 | 0.08 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | 0.41 | 0.06 | 0.08 | 0.03 | [9 empty] | 3
| 1 | 3 | 0.027 | 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | 0.41 | 0.37 | 0.26 | 0.02 | [9 empty] | 16
| 1 | 4 | 0.024 | 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | 0.43 | 0.06 | 0.2 | 0.03 | [9 empty] | 0.0442
| 1 | 5 | 0.024 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | 0.41 | 0.09 | 0.17 | 0.02 | [9 empty] | 26
| 1 | 6 | 0.024 | 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | 0.42 | 0.1 | 0.14 | 0.02
| Mean |  | 0.025167 | 0.05 | 0.08 | [5 empty] | 0.41833 | 0.255 | 0.155 | 0.025
| Median |  | 0.025 | 0.05 | 0.08 | [5 empty] | 0.415 | 0.095 | 0.155 | 0.025
| Std Dev. |  | 0.0013292 | 0 | 0 | [5 empty] | 0.0098319 | 0.3145 | 0.070356 | 0.0054772
| Rel.Std.Dev. |  | 5.281% | 0% | 0% | [5 empty] | 2.35% | 123.3% | 45.39% | 21.91%
| PDM3 |  | -43.06% | 13.12% | 81% | [5 empty] | 846.5% | 476.9% | 250.7% | -43.44%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 0.04 | 0.67 | 0.67 | 0.19 | 0.67 | 1.91 | 12.87 | 7.25 | 3.8 | 0.67
| Analytical results for Nd in OREAS 926b (Indicative Value 20.4 ppm) | [22 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 14 | 15 | 17 | 22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [16 empty] | 1
|  |  | 0.5g | 0.25g | 40g | 0.5g | 0.5g | [16 empty] | 1
| 1 | 1 | 24.8 | 25.5339 | 22.1 | 18 | 12.3 | [16 empty] | 1
| 1 | 2 | 25.7 | 25.5265 | 22 | 18.7 | 11.7 | [16 empty] | 4
| 1 | 3 | 23.2 | 25.2228 | 22 | 18.4 | 12.9 | [16 empty] | 16
| 1 | 4 | 24.7 | 25.3296 | 22.3 | 16.5 | 12.4 | [16 empty] | 20.4295
| 1 | 5 | 24 | 25.3191 | 22.4 | 16.3 | 12.5 | [16 empty] | 27
| 1 | 6 | 26.5 | 25.5939 | 21.2 | 16.7 | 12.1
| Mean |  | 24.8167 | 25.421 | 22 | 17.4333 | 12.3167
| Median |  | 24.75 | 25.428 | 22.05 | 17.35 | 12.35
| Std Dev. |  | 1.1754 | 0.14952 | 0.42426 | 1.0539 | 0.40208
| Rel.Std.Dev. |  | 4.737% | 0.5882% | 1.928% | 6.045% | 3.265%
| PDM3 |  | 21.47% | 24.43% | 7.687% | -14.67% | -39.71%
| Z-Score (Absolute) |  | 0.56 | 0.67 | 0 | 0.9 | 1.91
| Analytical results for Ni in OREAS 926b (Certified Value 6.9 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 6.27 | 6.5 | 7.2 | 6.6 | 6.8 | 7.2 | 10.9 | 7.1 | 7 | 6.54 | 7 | 7.2 | 7.1 | 5.6 | 7.118 | 6.2 | 7.5325 | 7.23 |  |  |  | 1
| 1 | 2 | 6.64 | 6.5 | 7 | 6.6 | 6.9 | 7.7 | 13.6 | 6.6 | 7 | 6.47 | 7 | 7.3 | 6.79 | 6 | 7.266 | 6.9 | 8.4765 | 6.76 |  |  |  | 5
| 1 | 3 | 6.43 | 6.5 | 6.6 | 6.8 | 7.2 | 7.2 | 8.3 | 7 | 7 | 6.41 | 6 | 7.4 | 6.95 | 5.2 | 8.308 | 6.5 | 9.7261 | 7.22 |  |  |  | 16
| 1 | 4 | 6.57 | 6.5 | 6.7 | 6.9 | 6.9 | 7.4 | 7.2 | 6.8 | 7 | 6.6 | 6 | 7.2 | 6.71 | 5.6 | 8.496 | 7.1 | 7.8819 | 7.16 |  |  |  | 6.8989
| 1 | 5 | 6.56 | 6.6 | 6.6 | 7.2 | 6.7 | 7.4 | 7.3 | 6.5 | 7 | 6.76 | 6 | 7.3 | 6.54 | 6 | 7.536 | 6.6 | 8.3774 | 7.15 |  |  |  | 161
| 1 | 6 | 7.3 | 6.7 | 6.3 | 7 | 6.5 | 7.4 | 6.9 | 7 | 7 | 6.71 | 6 | 7.2 | 6.42 | 6.5 | 7.838 | 7.2 | 8.3567 | 6.99
| Mean |  | 6.6283 | 6.55 | 6.7333 | 6.85 | 6.8333 | 7.3833 | 9.0333 | 6.8333 | 7 | 6.5817 | 6.3333 | 7.2667 | 6.7517 | 5.8167 | 7.7603 | 6.75 | 8.3918 | 7.085
| Median |  | 6.565 | 6.5 | 6.65 | 6.85 | 6.85 | 7.4 | 7.8 | 6.9 | 7 | 6.57 | 6 | 7.25 | 6.75 | 5.8 | 7.687 | 6.75 | 8.3671 | 7.155
| Std Dev. |  | 0.35414 | 0.083666 | 0.32042 | 0.23452 | 0.23381 | 0.18348 | 2.6756 | 0.24221 | 0 | 0.13586 | 0.5164 | 0.08165 | 0.25246 | 0.44907 | 0.55758 | 0.38341 | 0.74663 | 0.18097
| Rel.Std.Dev. |  | 5.343% | 1.277% | 4.759% | 3.424% | 3.422% | 2.485% | 29.62% | 3.545% | 0% | 2.064% | 8.154% | 1.124% | 3.739% | 7.72% | 7.185% | 5.68% | 8.897% | 2.554%
| PDM3 |  | -3.922% | -5.058% | -2.4% | -0.7093% | -0.9509% | 7.021% | 30.94% | -0.9509% | 1.465% | -4.599% | -8.198% | 5.33% | -2.135% | -15.69% | 12.49% | -2.159% | 21.64% | 2.697%
| Z-Score (Absolute) |  | 0.55 | 0.76 | 0.27 | 0.04 | 0 | 1.47 | 5.89 | 0 | NA | 0.67 | NA | 1.16 | 0.22 | 2.72 | 2.48 | 0.22 | 4.18 | 0.67
Results from laboratories 13 and 15 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 926b (Certified Value 0.021 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g |  | 0.5g |  |  |  | 3
| 1 | 1 | 0.02 | 0.021 | 0.021 | 0.02 | 0.02 | 0.025 | 0.04 | 0.021 | 0.0229 | 0.046392 | 0.022 | 0.021 | 0.0209 | 0.0275 | 0.026543 | 0.0203 | 0.0504 | < 0.0009 |  |  |  | 1
| 1 | 2 | 0.021 | 0.021 | 0.021 | 0.02 | 0.02 | 0.025 | 0.04 | 0.021 | 0.0226 | 0.046478 | 0.02 | 0.02 | 0.021 | 0.0189 | 0.027864 | 0.0199 | 0.0486 | < 0.0009 |  |  |  | #N/A
| 1 | 3 | 0.021 | 0.021 | 0.021 | 0.02 | 0.02 | 0.025 | 0.04 | 0.02 | 0.0228 | 0.046626 | 0.02 | 0.02 | 0.0212 | 0.0192 | 0.027686 | 0.0205 | 0.0486 | < 0.0009 |  |  |  | 16
| 1 | 4 | 0.021 | 0.021 | 0.021 | 0.02 | 0.02 | 0.026 | 0.039 | 0.02 | 0.0227 | 0.046463 | 0.02 | 0.02 | 0.0203 | 0.0201 | 0.026536 | 0.0217 | 0.0504 | < 0.0009 |  |  |  | 0.020526
| 1 | 5 | 0.02 | 0.021 | 0.021 | 0.02 | 0.019 | 0.027 | 0.037 | 0.02 | 0.0228 | 0.046522 | 0.02 | 0.02 | 0.0199 | 0.018 | 0.026301 | 0.0215 | 0.0504 | < 0.0009 |  |  |  | 162
| 1 | 6 | 0.021 | 0.021 | 0.021 | 0.02 | 0.02 | 0.025 | 0.038 | 0.02 | 0.0227 | 0.046502 | 0.02 | 0.02 | 0.02 | 0.0238 | 0.02881 | 0.0199 | 0.0472 | < 0.0009
| Mean |  | 0.020667 | 0.021 | 0.021 | 0.02 | 0.019833 | 0.0255 | 0.039 | 0.020333 | 0.02275 | 0.046497 | 0.020333 | 0.020167 | 0.02055 | 0.02125 | 0.02729 | 0.020633 | 0.049267
| Median |  | 0.021 | 0.021 | 0.021 | 0.02 | 0.02 | 0.025 | 0.0395 | 0.02 | 0.02275 | 0.04649 | 0.02 | 0.02 | 0.0206 | 0.01965 | 0.027115 | 0.0204 | 0.0495
| Std Dev. |  | 0.0005164 | 0 | 0 | 0 | 0.00040825 | 0.00083666 | 0.0012649 | 0.0005164 | 0.00010488 | 7.7157e-05 | 0.0008165 | 0.00040825 | 0.00055408 | 0.0036654 | 0.00099012 | 0.00078655 | 0.0013426
| Rel.Std.Dev. |  | 2.499% | 0% | 0% | 0% | 2.058% | 3.281% | 3.243% | 2.54% | 0.461% | 0.1659% | 4.016% | 2.024% | 2.696% | 17.25% | 3.628% | 3.812% | 2.725%
| PDM3 |  | 0.6834% | 2.307% | 2.307% | -2.564% | -3.376% | 24.23% | 90% | -0.9405% | 10.83% | 126.5% | -0.9405% | -1.752% | 0.115% | 3.525% | 32.95% | 0.521% | 140%
| Z-Score (Absolute) |  | 0.15 | 0.15 | 0.15 | 0.75 | 0.9 | 4.2 | 16.33 | 0.45 | 1.72 | 23.07 | 0.45 | 0.6 | 0.25 | 0.37 | 5.81 | 0.18 | 25.56 | 18.34
| Analytical results for Pb in OREAS 926b (Certified Value 40.3 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 39 | 42.6 | 36 | 37.4 | 39.9 | 43.6 | 140 | 35.81 | 105 | 45.8 | 1251.015 | 44 | 44.1 | 45 | 54.4 | 47.799 | 45.1 | 42.5063 | 49.1 |  |  | 1
| 1 | 2 | 36.4 | 40.6 | 37.9 | 37 | 38.8 | 53.6 | 66.2 | 34.54 | 47 | 44.5 | 128.0925 | 43 | 40.7 | 46 | 40.2 | 40.257 | 34.2 | 36.8083 | 56 |  |  | 6
| 1 | 3 | 36.3 | 40 | 36.9 | 38.8 | 37.6 | 44.1 | 58.3 | 36.9 | 43 | 42.8 | 61.575 | 42 | 39.6 | 49 | 33.8 | 41.372 | 32 | 37.1284 | 39.6 |  |  | 16
| 1 | 4 | 35.4 | 40.3 | 37.4 | 38 | 37.9 | 41.2 | 38.1 | 32.64 | 43 | 43.3 | 84.2235 | 40 | 40.3 | 44 | 36.7 | 40.159 | 32.4 | 38.2193 | 36.3 |  |  | 40.3209
| 1 | 5 | 35.3 | 39.6 | 36.8 | 38.7 | 37 | 41.2 | 34.5 | 32.61 | 40 | 42.3 | 43.83 | 40 | 40.1 | 44 | 36.1 | 37.534 | 31.5 | 36.5307 | 45 |  |  | 163
| 1 | 6 | 36.9 | 40.6 | 34.9 | 37.7 | 36.9 | 43.1 | 32.7 | 33.36 | 42 | 43.2 | 59.5065 | 40 | 40.8 | 44 | 36.7 | 38.202 | 30.9 | 38.2255 | 36.4
| Mean |  | 36.55 | 40.6167 | 36.65 | 37.9333 | 38.0167 | 44.4667 | 61.6333 | 34.31 | 53.3333 | 43.65 | 271.3738 | 41.5 | 40.9333 | 45.3333 | 39.65 | 40.8872 | 34.35 | 38.2364 | 43.7333
| Median |  | 36.35 | 40.45 | 36.85 | 37.85 | 37.75 | 43.35 | 48.2 | 33.95 | 43 | 43.25 | 72.8992 | 41 | 40.5 | 44.5 | 36.7 | 40.208 | 32.2 | 37.6739 | 42.3
| Std Dev. |  | 1.3487 | 1.0439 | 1.0672 | 0.71461 | 1.1514 | 4.6371 | 40.7443 | 1.7659 | 25.4139 | 1.2818 | 480.8183 | 1.7607 | 1.6108 | 1.9664 | 7.5113 | 3.6712 | 5.384 | 2.2095 | 7.8357
| Rel.Std.Dev. |  | 3.69% | 2.57% | 2.912% | 1.884% | 3.029% | 10.43% | 66.11% | 5.147% | 47.65% | 2.937% | 177.2% | 4.243% | 3.935% | 4.338% | 18.94% | 8.979% | 15.67% | 5.778% | 17.92%
| PDM3 |  | -9.352% | 0.7334% | -9.104% | -5.922% | -5.715% | 10.28% | 52.86% | -14.91% | 32.27% | 8.256% | 573% | 2.924% | 1.519% | 12.43% | -1.664% | 1.404% | -14.81% | -5.17% | 8.463%
| Z-Score (Absolute) |  | 0.99 | 0.06 | 0.97 | 0.67 | 0.66 | 0.82 | 4.74 | 1.5 | 2.84 | 0.63 | 52.62 | 0.14 | 0.01 | 1.01 | 0.28 | 0 | 1.49 | 0.61 | 0.65
| Analytical results for Pd in OREAS 926b (Indicative Value 31.3 ppb) | [22 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 03 | 15 | [19 empty] | ppb
|  |  | AR*MS | AR*MS | [19 empty] | 1
|  |  | 0.5g | 40g | [19 empty] | 1
| 1 | 1 | 2 | 60 | [19 empty] | 1
| 1 | 2 | 3 | 50 | [19 empty] | 22
| 1 | 3 | 6 | 60 | [19 empty] | 16
| 1 | 4 | 4 | 60 | [19 empty] | 31.25
| 1 | 5 | 6 | 60 | [19 empty] | 28
| 1 | 6 | 4 | 60
| Mean |  | 4.1667 | 58.3333
| Median |  | 4 | 60
| Std Dev. |  | 1.6021 | 4.0825
| Rel.Std.Dev. |  | 38.45% | 6.999%
| PDM3 |  | -86.67% | 86.67%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 926b (Indicative Value 5.55 ppm) | [22 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 14 | 15 | 17 | 22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [16 empty] | 2
|  |  | 0.5g | 0.25g | 40g | 0.5g | 0.5g | [16 empty] | 2
| 1 | 1 | 6.6 | 5.3851 | 5.41 | 5.07 | 3.42 | [16 empty] | 1
| 1 | 2 | 7 | 5.4515 | 5.4 | 5.07 | 3.27 | [16 empty] | 7
| 1 | 3 | 6.1 | 5.3243 | 5.33 | 4.97 | 3.61 | [16 empty] | 16
| 1 | 4 | 6.4 | 5.2012 | 5.45 | 4.85 | 3.48 | [16 empty] | 5.5529
| 1 | 5 | 6.4 | 5.1511 | 5.48 | 4.76 | 3.54 | [16 empty] | 29
| 1 | 6 | 7 | 5.2332 | 5.17 | 4.82 | 3.46
| Mean |  | 6.5833 | 5.2911 | 5.3733 | 4.9233 | 3.4633
| Median |  | 6.5 | 5.2788 | 5.405 | 4.91 | 3.47
| Std Dev. |  | 0.36009 | 0.11549 | 0.11183 | 0.13261 | 0.1157
| Rel.Std.Dev. |  | 5.47% | 2.183% | 2.081% | 2.694% | 3.341%
| PDM3 |  | 18.56% | -4.716% | -3.234% | -11.34% | -37.63%
| Z-Score (Absolute) |  | 2.37 | 0 | 0.15 | 0.67 | 3.35
| Analytical results for Pt in OREAS 926b (Indicative Value < 10 ppb) | [22 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [19 empty] | 1
| No | No | 03 | 15 | [19 empty] | ppb
|  |  | AR*MS | AR*MS | [19 empty] | 1
|  |  | 0.5g | 40g | [19 empty] | 1
| 1 | 1 | 2 | < 10 | [19 empty] | 1
| 1 | 2 | < 2 | < 10 | [19 empty] | 24
| 1 | 3 | < 2 | < 10 | [19 empty] | 16
| 1 | 4 | < 2 | < 10 | [19 empty] | < 10
| 1 | 5 | < 2 | < 10 | [19 empty] | 30
| 1 | 6 | 2 | < 10
| Mean |  | 2
| Median |  | 2
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 926b (Certified Value 10.2 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 15 | 17 | 18 | 20 | 21 | 22 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 10.1 | 11.2 | 11.4 | 11 | 10.2 | 13.4 | 9.3 | 16 | 11.6 | 11.43 | 7.3 | 8.1 | 17.5804 | 8.08 | [7 empty] | 1
| 1 | 2 | 10.55 | 10.6 | 11.4 | 10.8 | 10.6 | 13.5 | 9.3 | 16.1 | 11.5 | 11.37 | 8.6 | 8.3 | 17.3976 | 7.89 | [7 empty] | 8
| 1 | 3 | 10.05 | 10.7 | 10.6 | 10.4 | 10.3 | 12.6 | 9.4 | 15.9 | 11.6 | 10.72 | 7.6 | 8.4 | 16.9374 | 8.48 | [7 empty] | 16
| 1 | 4 | 10.1 | 10.8 | 10.9 | 10.2 | 10.1 | 13.8 | 9.5 | 16.3 | 11.7 | 9.95 | 8.4 | 8.5 | 17.6901 | 8.21 | [7 empty] | 10.1878
| 1 | 5 | 10.25 | 10.8 | 10.7 | 11 | 9.7 | 13.6 | 9.4 | 16.3 | 11.9 | 9.56 | 8.4 | 8.2 | 17.8892 | 8.46 | [7 empty] | 164
| 1 | 6 | 10.6 | 11.2 | 10.2 | 10.7 | 9.8 | 14.2 | 8.9 | 16.4 | 11.2 | 9.27 | 8.6 | 8.2 | 17.2093 | 8.15
| Mean |  | 10.275 | 10.8833 | 10.8667 | 10.6833 | 10.1167 | 13.5167 | 9.3 | 16.1667 | 11.5833 | 10.3833 | 8.15 | 8.2833 | 17.4507 | 8.2117
| Median |  | 10.175 | 10.8 | 10.8 | 10.75 | 10.15 | 13.55 | 9.35 | 16.2 | 11.6 | 10.335 | 8.4 | 8.25 | 17.489 | 8.18
| Std Dev. |  | 0.24238 | 0.25626 | 0.47188 | 0.32506 | 0.33116 | 0.53072 | 0.20976 | 0.19664 | 0.23166 | 0.92602 | 0.55767 | 0.1472 | 0.34381 | 0.22728
| Rel.Std.Dev. |  | 2.359% | 2.355% | 4.342% | 3.043% | 3.273% | 3.926% | 2.256% | 1.216% | 2% | 8.918% | 6.843% | 1.777% | 1.97% | 2.768%
| PDM3 |  | 0.8561% | 6.827% | 6.664% | 4.864% | -0.698% | 32.68% | -8.714% | 58.69% | 13.7% | 1.92% | -20% | -18.69% | 71.29% | -19.4%
| Z-Score (Absolute) |  | 0.15 | 0.21 | 0.2 | 0.09 | 0.25 | 1.76 | 0.73 | 3.33 | 0.62 | 0.09 | 1.41 | 1.33 | 4.09 | 1.37
| Analytical results for Re in OREAS 926b (Indicative Value 0.883 ppb) | [22 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 17 | 18 | 20 | [10 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.5 | 1 | 1 | 1 | 2 | < 1 | 1 | < 50 | 3 | < 100 | 1 | [10 empty] | 1
| 1 | 2 | 0.6 | 1 | < 1 | 1 | 1 | < 1 | < 1 | < 50 | 3 | < 100 | 1 | [10 empty] | 12
| 1 | 3 | 0.6 | 1 | 1 | 1 | 1 | < 1 | 1 | < 50 | 3 | < 100 | 1 | [10 empty] | 16
| 1 | 4 | 0.5 | 1 | 1 | < 1 | 1 | < 1 | 2 | < 50 | 3 | < 100 | < 1 | [10 empty] | 0.88333
| 1 | 5 | 0.6 | 1 | 1 | 1 | 1 | < 1 | 1 | < 50 | 3 | < 100 | < 1 | [10 empty] | 25
| 1 | 6 | 0.4 | < 1 | < 1 | 1 | 1 | < 1 | < 1 | < 50 | 3 | < 100 | 1
| Mean |  | 0.53333 | 1 | 1 | 1 | 1.1667 |  | 1.25 |  | 3 |  | 1
| Median |  | 0.55 | 1 | 1 | 1 | 1 |  | 1 |  | 3 |  | 1
| Std Dev. |  | 0.08165 | 0 | 0 | 0 | 0.40825 |  | 0.5 |  | 0 |  | 0
| Rel.Std.Dev. |  | 15.31% | 0% | 0% | 0% | 34.99% |  | 40% |  | 0% |  | 0%
| PDM3 |  | -39.62% | 13.21% | 13.21% | 13.21% | 32.08% |  | 41.51% |  | 239.6% |  | 13.21%
| Z-Score (Absolute) |  | 1.03 | 0 | 0.22 | 0 | 0.67 | 1.12 | 0.22 | 64.96 | 5.62 | 132.39 | 0.22
| Analytical results for S in OREAS 926b (Certified Value 0.993 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.98 | 1.03 | 1.02 | 0.96 | 0.97 | 0.994 | 1.13 | 0.98 | 1.0053 | 0.95 | 0.95584 | 1.03 | 1 | 0.94 | 0.92 | 0.873 | 0.87 | 0.957 | 0.998 |  |  | 1
| 1 | 2 | 1 | 1.04 | 1.02 | 0.96 | 0.99 | 1 | 1.08 | 0.96 | 1.0102 | 0.95 | 0.95575 | 1.03 | 1.02 | 0.95 | 1.05 | 0.849 | 0.92 | 0.979 | 0.979 |  |  | #N/A
| 1 | 3 | 1 | 1.04 | 1.05 | 1 | 0.98 | 1.007 | 1.08 | 0.96 | 1.0187 | 0.95 | 0.95264 | 0.98 | 1.02 | 0.99 | 0.89 | 0.867 | 0.98 | 0.963 | 1.01 |  |  | 16
| 1 | 4 | 0.99 | 1.04 | 1.05 | 1.04 | 0.97 | 1.011 | 1.09 | 0.97 | 1.0222 | 0.96 | 0.95845 | 1.02 | 1.02 | 0.96 | 1.01 | 0.864 | 0.98 | 0.939 | 0.981 |  |  | 0.99324
| 1 | 5 | 0.99 | 1.04 | 1.04 | 1 | 0.96 | 1.051 | 1.1 | 0.96 | 1.0043 | 0.94 | 0.95473 | 0.97 | 1.03 | 0.97 | 1.04 | 0.869 | 0.96 | 0.95 | 0.988 |  |  | 165
| 1 | 6 | 1.01 | 1.05 | 1.02 | 1.01 | 0.95 | 1 | 1.04 | 0.97 | 1.0254 | 0.95 | 0.9519 | 0.99 | 1.02 | 0.92 | 1.06 | 0.855 | 0.95 | 0.997 | 0.993
| Mean |  | 0.995 | 1.04 | 1.0333 | 0.995 | 0.97 | 1.0105 | 1.0867 | 0.96667 | 1.0144 | 0.95 | 0.95489 | 1.0033 | 1.0183 | 0.955 | 0.995 | 0.86283 | 0.94333 | 0.96417 | 0.9915
| Median |  | 0.995 | 1.04 | 1.03 | 1 | 0.97 | 1.0035 | 1.085 | 0.965 | 1.0145 | 0.95 | 0.95524 | 1.005 | 1.02 | 0.955 | 1.025 | 0.8655 | 0.955 | 0.96 | 0.9905
| Std Dev. |  | 0.010488 | 0.0063246 | 0.015055 | 0.030822 | 0.014142 | 0.020715 | 0.029439 | 0.008165 | 0.0089752 | 0.0063246 | 0.0023808 | 0.026583 | 0.0098319 | 0.02429 | 0.072319 | 0.0090866 | 0.042269 | 0.020904 | 0.011537
| Rel.Std.Dev. |  | 1.054% | 0.6081% | 1.457% | 3.098% | 1.458% | 2.05% | 2.709% | 0.8447% | 0.8848% | 0.6657% | 0.2493% | 2.649% | 0.9655% | 2.543% | 7.268% | 1.053% | 4.481% | 2.168% | 1.164%
| PDM3 |  | 0.1767% | 4.707% | 4.036% | 0.1767% | -2.34% | 1.737% | 9.406% | -2.676% | 2.125% | -4.354% | -3.862% | 1.016% | 2.526% | -3.85% | 0.1767% | -13.13% | -5.025% | -2.928% | -0.1756%
| Z-Score (Absolute) |  | 0 | 1.07 | 0.91 | 0 | 0.59 | 0.37 | 2.18 | 0.67 | 0.46 | 1.07 | 0.95 | 0.2 | 0.56 | 0.95 | 0 | 3.15 | 1.23 | 0.73 | 0.08
| Analytical results for Sb in OREAS 926b (Certified Value 3.03 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 2.28 | 2.88 | 2.36 | 3.02 | 2.43 | 1.28 | 4.35 | 2.91 | 3.27 | 3.68 | 2.4658 | 11.8 | 3.5 | 3.09 | 6.41 | < 10 | 3.41 | 2.5396 | 3.27 |  |  | 1
| 1 | 2 | 2.29 | 2.93 | 2.32 | 2.98 | 2.5 | 1.12 | 2.17 | 2.83 | 3.24 | 3.64 | 2.624 | 6.1 | 3.7 | 3.13 | 3.48 | < 10 | 3.2 | 2.229 | 3.72 |  |  | 9
| 1 | 3 | 2.22 | 2.87 | 2.37 | 2.63 | 2.41 | 1.49 | 4.02 | 2.76 | 3.39 | 3.7 | 2.7441 | 4.3 | 3.6 | 3.18 | 3.16 | < 10 | 3.18 | 2.4024 | 3.98 |  |  | 16
| 1 | 4 | 2.27 | 2.9 | 2.46 | 2.6 | 2.49 | 1.69 | 4.42 | 2.72 | 3.36 | 3.57 | 2.524 | 3.4 | 3.5 | 3.14 | 3.45 | < 10 | 3.16 | 2.3253 | 3.66 |  |  | 3.0309
| 1 | 5 | 2.25 | 2.89 | 2.38 | 2.71 | 2.4 | 1.27 | 2.04 | 2.62 | 3.16 | 3.59 | 2.5836 | 3.3 | 3.5 | 3.23 | 3.42 | < 10 | 3.02 | 2.3452 | 4.15 |  |  | 166
| 1 | 6 | 2.25 | 2.94 | 2.37 | 2.58 | 2.44 | 1.22 | 4.75 | 2.77 | 3.17 | 3.66 | 2.4718 | 3.1 | 3.6 | 2.99 | 3.44 | < 10 | 3.04 | 2.3677 | 3.88
| Mean |  | 2.26 | 2.9017 | 2.3767 | 2.7533 | 2.445 | 1.345 | 3.625 | 2.7683 | 3.265 | 3.64 | 2.5689 | 5.3333 | 3.5667 | 3.1267 | 3.8933 |  | 3.1683 | 2.3682 | 3.7767
| Median |  | 2.26 | 2.895 | 2.37 | 2.67 | 2.435 | 1.275 | 4.185 | 2.765 | 3.255 | 3.65 | 2.5538 | 3.85 | 3.55 | 3.135 | 3.445 |  | 3.17 | 2.3564 | 3.8
| Std Dev. |  | 0.025298 | 0.027869 | 0.045898 | 0.19654 | 0.041352 | 0.20792 | 1.2007 | 0.098268 | 0.095237 | 0.05099 | 0.10585 | 3.3566 | 0.08165 | 0.082138 | 1.2384 |  | 0.14006 | 0.10225 | 0.30494
| Rel.Std.Dev. |  | 1.119% | 0.9604% | 1.931% | 7.138% | 1.691% | 15.46% | 33.12% | 3.55% | 2.917% | 1.401% | 4.12% | 62.94% | 2.289% | 2.627% | 31.81% |  | 4.421% | 4.318% | 8.074%
| PDM3 |  | -25.43% | -4.264% | -21.59% | -9.158% | -19.33% | -55.62% | 19.6% | -8.663% | 7.724% | 20.1% | -15.24% | 75.96% | 17.68% | 3.159% | 28.45% |  | 4.534% | -21.86% | 24.61%
| Z-Score (Absolute) |  | 1.05 | 0.27 | 0.91 | 0.45 | 0.82 | 2.15 | 0.6 | 0.43 | 0.17 | 0.62 | 0.67 | 2.67 | 0.53 | 0 | 0.93 | 2.26 | 0.05 | 0.92 | 0.79
| Analytical results for Sc in OREAS 926b (Certified Value 1.27 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 1.215 | 1.2 | 1.2 | 1.3 | 1.3 | 1.7 | 1.5 | 1.3 | 1 | 1.3278 | 1.5 | 1.1 | 1.4 | 1.1 | 1.1 | 1.3043 | 1 |  |  |  |  | 1
| 1 | 2 | 1.285 | 1.2 | 1.3 | 1.2 | 1.3 | 1.6 | 1.4 | 1.2 | 1 | 1.2719 | 1.5 | 1.2 | 1.3 | 1.1 | 1 | 1.1789 | 0.9 |  |  |  |  | 13
| 1 | 3 | 1.205 | 1.2 | 1.3 | 1.3 | 1.3 | 1.5 | 1.4 | 1.3 | 1 | 1.2187 | 1.5 | 1.2 | 1.3 | 1 | 1.3 | 1.5819 | 0.9 |  |  |  |  | 16
| 1 | 4 | 1.225 | 1.2 | 1.3 | 1.3 | 1.3 | 1.7 | 1.4 | 1.2 | 1 | 1.3713 | 1.5 | 1.2 | 1.2 | 1 | 1.1 | 1.4964 | 0.9 |  |  |  |  | 1.2727
| 1 | 5 | 1.27 | 1.2 | 1.3 | 1.4 | 1.3 | 1.7 | 1.4 | 1.2 | 1 | 1.2813 | 1.5 | 1.2 | 1.3 | 1.1 | 1 | 1.5672 | 0.9 |  |  |  |  | 167
| 1 | 6 | 1.24 | 1.2 | 1.2 | 1.3 | 1.3 | 1.8 | 1.4 | 1.2 | 1 | 1.4172 | 1.5 | 1.2 | 1.2 | 1.1 | 1 | 1.4464 | 0.9
| Mean |  | 1.24 | 1.2 | 1.2667 | 1.3 | 1.3 | 1.6667 | 1.4167 | 1.2333 | 1 | 1.3147 | 1.5 | 1.1833 | 1.2833 | 1.0667 | 1.0833 | 1.4292 | 0.91667
| Median |  | 1.2325 | 1.2 | 1.3 | 1.3 | 1.3 | 1.7 | 1.4 | 1.2 | 1 | 1.3046 | 1.5 | 1.2 | 1.3 | 1.1 | 1.05 | 1.4714 | 0.9
| Std Dev. |  | 0.031623 | 0 | 0.05164 | 0.063246 | 0 | 0.10328 | 0.040825 | 0.05164 | 0 | 0.072176 | 0 | 0.040825 | 0.075277 | 0.05164 | 0.1169 | 0.15838 | 0.040825
| Rel.Std.Dev. |  | 2.55% | 0% | 4.077% | 4.865% | 0% | 6.197% | 2.882% | 4.187% | 0% | 5.49% | 0% | 3.45% | 5.866% | 4.841% | 10.79% | 11.08% | 4.454%
| PDM3 |  | -2.566% | -5.709% | -0.4709% | 2.148% | 2.148% | 30.96% | 11.32% | -3.09% | -21.42% | 3.304% | 17.86% | -7.019% | 0.8387% | -16.19% | -14.88% | 12.3% | -27.97%
| Z-Score (Absolute) |  | 0.28 | 0.61 | 0.07 | 0.2 | 0.2 | 3.17 | 1.15 | 0.34 | NA | 0.32 | 1.82 | 0.74 | 0.07 | 1.69 | 1.55 | 1.25 | 2.9
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 926b (Certified Value 11.4 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 11.65 | 10.9 | 11.4 | 11.4 | 11.4 | 12.3 | 7.7 | 11.2 | 10.24 | 11 | 8 | 12 | 9.4 | 13 | 11.4 | 13.5241 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 11.85 | 11.6 | 11.7 | 11.4 | 12.3 | 12.9 | 6.9 | 10.9 | 10.41 | 11 | 8.5 | 12 | 9.46 | 14 | 11 | 15.8865 | < 0.1 |  |  |  |  | 14
| 1 | 3 | 11.9 | 10.6 | 11.5 | 11.4 | 12.2 | 11.5 | 7.2 | 11.2 | 11.14 | 11 | 8.5 | 12 | 8.9 | 12 | 11.1 | 15.6468 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 11.7 | 11 | 12.4 | 11.5 | 11.5 | 12.6 | 9.4 | 11.3 | 11.07 | 11 | 8.5 | 12 | 10.09 | 14 | 11.5 | 15.6801 | < 0.1 |  |  |  |  | 11.3764
| 1 | 5 | 11.5 | 10.8 | 11.6 | 11.7 | 10.6 | 11.8 | 7.8 | 10.7 | 9.96 | 11 | 8.5 | 12 | 8.82 | 14 | 11 | 16.3136 | < 0.1 |  |  |  |  | 168
| 1 | 6 | 11.9 | 10.8 | 11.1 | 11.7 | 10.8 | 12.8 | 6.8 | 11 | 11.03 | 11 | 8.5 | 11 | 8.7 | 14 | 11.4 | 15.4623 | < 0.1
| Mean |  | 11.75 | 10.95 | 11.6167 | 11.5167 | 11.4667 | 12.3167 | 7.6333 | 11.05 | 10.6417 | 11 | 8.4167 | 11.8333 | 9.2283 | 13.5 | 11.2333 | 15.4189
| Median |  | 11.775 | 10.85 | 11.55 | 11.45 | 11.45 | 12.45 | 7.45 | 11.1 | 10.72 | 11 | 8.5 | 12 | 9.15 | 14 | 11.25 | 15.6635
| Std Dev. |  | 0.16125 | 0.34496 | 0.43551 | 0.1472 | 0.69761 | 0.56362 | 0.95638 | 0.22583 | 0.50245 | 0 | 0.20412 | 0.40825 | 0.52522 | 0.83666 | 0.22509 | 0.97276
| Rel.Std.Dev. |  | 1.372% | 3.15% | 3.749% | 1.278% | 6.084% | 4.576% | 12.53% | 2.044% | 4.722% | 0% | 2.425% | 3.45% | 5.691% | 6.197% | 2.004% | 6.309%
| PDM3 |  | 3.284% | -3.748% | 2.112% | 1.233% | 0.7938% | 8.265% | -32.9% | -2.869% | -6.458% | -3.308% | -26.02% | 4.017% | -18.88% | 18.67% | -1.257% | 35.53%
| Z-Score (Absolute) |  | 0.59 | 0.32 | 0.44 | 0.32 | 0.27 | 1.23 | 4.1 | 0.21 | 0.67 | NA | 3.21 | NA | 2.29 | 2.58 | 0 | 4.77 | 12.75
Results from laboratories 13 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 926b (Certified Value 3.62 ppm) | [22 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 14 | 15 | 17 | 22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [16 empty] | 2
|  |  | 0.5g | 0.25g | 40g | 0.5g | 0.5g | [16 empty] | 2
| 1 | 1 | 3.6 | 4.4838 | 3.68 | 3.49 | 2.45 | [16 empty] | 1
| 1 | 2 | 4.6 | 4.4357 | 3.57 | 3.66 | 2.38 | [16 empty] | 1
| 1 | 3 | 4.3 | 4.3068 | 3.53 | 3.58 | 2.6 | [16 empty] | 16
| 1 | 4 | 4.4 | 4.3075 | 3.65 | 3.23 | 2.5 | [16 empty] | 3.6185
| 1 | 5 | 3.8 | 4.4574 | 3.64 | 3.38 | 2.62 | [16 empty] | 169
| 1 | 6 | 4.5 | 4.2546 | 3.44 | 3.3 | 2.41
| Mean |  | 4.2 | 4.3743 | 3.585 | 3.44 | 2.4933
| Median |  | 4.35 | 4.3716 | 3.605 | 3.435 | 2.475
| Std Dev. |  | 0.40497 | 0.095901 | 0.090056 | 0.16601 | 0.09913
| Rel.Std.Dev. |  | 9.642% | 2.192% | 2.512% | 4.826% | 3.976%
| PDM3 |  | 16.07% | 20.89% | -0.9266% | -4.934% | -31.1%
| Z-Score (Absolute) |  | 0.67 | 0.87 | 0 | 0.16 | 1.2
| Analytical results for Sn in OREAS 926b (Certified Value 7.5 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 15 | 17 | 18 | 19 | 20 | 22 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 7.26 | 7.6 | 7.8 | 7.2 | 7.4 | 7.9 | 4.8 | 7.5 | 9 | 7.52 | 8.5 | < 10 | 6.9 | 7.67 | [7 empty] | 1
| 1 | 2 | 7.57 | 7.7 | 7.7 | 7.1 | 7.6 | 8.13 | 4.6 | 7.5 | 9 | 7.43 | 9.5 | < 10 | 7.1 | 7.4 | [7 empty] | 2
| 1 | 3 | 7.38 | 7.7 | 7.8 | 7.2 | 7.6 | 7.35 | 4.5 | 7.5 | 9 | 7.6 | 8.1 | < 10 | 7.1 | 7.85 | [7 empty] | 16
| 1 | 4 | 7.6 | 7.7 | 7.9 | 7.3 | 7.6 | 7.81 | 4.6 | 7.6 | 9 | 7.43 | 9.4 | < 10 | 7.1 | 7.49 | [7 empty] | 7.5017
| 1 | 5 | 7.46 | 7.6 | 7.7 | 7.4 | 7.3 | 7.68 | 4.5 | 7.4 | 9 | 7.4 | 9.3 | < 10 | 6.8 | 7.61 | [7 empty] | 170
| 1 | 6 | 7.76 | 7.7 | 7.5 | 7.4 | 7.4 | 8.05 | 4.5 | 7.6 | 8 | 7.27 | 9.4 | < 10 | 6.9 | 7.58
| Mean |  | 7.505 | 7.6667 | 7.7333 | 7.2667 | 7.4833 | 7.82 | 4.5833 | 7.5167 | 8.8333 | 7.4417 | 9.0333 |  | 6.9833 | 7.6
| Median |  | 7.515 | 7.7 | 7.75 | 7.25 | 7.5 | 7.855 | 4.55 | 7.5 | 9 | 7.43 | 9.35 |  | 7 | 7.595
| Std Dev. |  | 0.17661 | 0.05164 | 0.13663 | 0.12111 | 0.13292 | 0.28142 | 0.1169 | 0.075277 | 0.40825 | 0.11197 | 0.58538 |  | 0.13292 | 0.15492
| Rel.Std.Dev. |  | 2.353% | 0.6736% | 1.767% | 1.667% | 1.776% | 3.599% | 2.551% | 1.001% | 4.622% | 1.505% | 6.48% |  | 1.903% | 2.038%
| PDM3 |  | 0.04443% | 2.2% | 3.088% | -3.133% | -0.2444% | 4.244% | -38.9% | 0.2% | 17.75% | -0.7998% | 20.42% |  | -6.91% | 1.311%
| Z-Score (Absolute) |  | 0 | 0.48 | 0.67 | 0.7 | 0.06 | 0.93 | 8.63 | 0.03 | NA | 0.19 | 4.51 | 7.4 | 1.54 | 0.28
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 926b (Indicative Value 8.15 ppm) | [22 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 7.69 | 8 | 9.4 | 8.1 | 7.7 | 12.6 | 5.3 | 7.5 | 10.6 | 12.78 | 8 | 8 | 8.5 | 3.3 | 9.85 | 3.9 | 9.6344 | 3 |  |  |  | 1
| 1 | 2 | 7.91 | 7.6 | 9.5 | 7.8 | 8.2 | 13.5 | 5.2 | 7.6 | 10.3 | 12.43 | 8 | 8 | 8.28 | 3.9 | 9.96 | 4 | 9.4197 | 4 |  |  |  | 3
| 1 | 3 | 7.69 | 7.4 | 8 | 7.2 | 7.8 | 11.6 | 5.2 | 7 | 10.1 | 12.69 | 8 | 8 | 8.28 | 3.3 | 10 | 4 | 9.1638 | 4 |  |  |  | 16
| 1 | 4 | 7.68 | 7.3 | 8.2 | 7.3 | 7.9 | 12.3 | 5.1 | 6.1 | 10.2 | 12.59 | 8 | 8 | 7.78 | 3.8 | 10.14 | 4.1 | 9.5188 | 4 |  |  |  | 8.1509
| 1 | 5 | 7.85 | 7.3 | 7.9 | 7.7 | 7.3 | 12.4 | 5.3 | 6.3 | 10.6 | 12.67 | 8 | 8 | 7.66 | 3.9 | 10.35 | 3.9 | 9.6905 | 4 |  |  |  | 26
| 1 | 6 | 8.09 | 7.6 | 7.7 | 7.7 | 7.2 | 13.9 | 5 | 7.4 | 10.7 | 12.89 | 7 | 8 | 7.54 | 4 | 9.94 | 3.8 | 9.4101 | 3
| Mean |  | 7.8183 | 7.5333 | 8.45 | 7.6333 | 7.6833 | 12.7167 | 5.1833 | 6.9833 | 10.4167 | 12.675 | 7.8333 | 8 | 8.0067 | 3.7 | 10.04 | 3.95 | 9.4729 | 3.6667
| Median |  | 7.77 | 7.5 | 8.1 | 7.7 | 7.75 | 12.5 | 5.2 | 7.2 | 10.45 | 12.68 | 8 | 8 | 8.03 | 3.85 | 9.98 | 3.95 | 9.4693 | 4
| Std Dev. |  | 0.16449 | 0.26583 | 0.79183 | 0.33267 | 0.37639 | 0.84242 | 0.1169 | 0.64317 | 0.24833 | 0.15796 | 0.40825 | 0 | 0.39551 | 0.31623 | 0.179 | 0.10488 | 0.18861 | 0.5164
| Rel.Std.Dev. |  | 2.104% | 3.529% | 9.371% | 4.358% | 4.899% | 6.625% | 2.255% | 9.21% | 2.384% | 1.246% | 5.212% | 0% | 4.94% | 8.547% | 1.783% | 2.655% | 1.991% | 14.08%
| PDM3 |  | -4.08% | -7.576% | 3.67% | -6.349% | -5.736% | 56.02% | -36.41% | -14.32% | 27.8% | 55.51% | -3.896% | -1.851% | -1.769% | -54.61% | 23.18% | -51.54% | 16.22% | -55.01%
| Z-Score (Absolute) |  | 0 | 0.12 | 0.26 | 0.08 | 0.06 | 2 | 1.07 | 0.34 | 1.06 | 1.98 | NA | NA | 0.08 | 1.68 | 0.91 | 1.58 | 0.67 | NA
Results from laboratories 15, 16 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 926b (Certified Value < 0.01 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 17 | 18 | 20 | 21 | 22 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [8 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | < 0.005 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 0.005 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | [8 empty] | 4
| 1 | 3 | < 0.005 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.05 | 0.01 | < 0.1 | < 0.01 | [8 empty] | 16
| 1 | 4 | < 0.005 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | [8 empty] | < 0.01
| 1 | 5 | < 0.005 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | [8 empty] | 171
| 1 | 6 | < 0.005 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | < 0.1 | < 0.01
| Mean |  |  |  | 0.01 | [5 empty] | 0.03 |  | 0.01
| Median |  |  |  | 0.01 | [5 empty] | 0.03 |  | 0.01
| Std Dev. |  |  |  | 0 | [5 empty] | 0
| Rel.Std.Dev. |  |  |  | 0% | [5 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 0.9 | 0.22 | 0.67 | 0.22 | 0.22 | 5.17 | 0.22 | 5.17 | 6.52 | 5.17 | 0 | 11.91 | 0.22
| Analytical results for Tb in OREAS 926b (Indicative Value 0.29 ppm) | [22 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 08 | 15 | 17 | 18 | 22 | [16 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.5g | [16 empty] | 2
| 1 | 1 | 0.4 | 0.345 | 0.3 | 0.18 | 0.21 | [16 empty] | 1
| 1 | 2 | 0.4 | 0.355 | 0.3 | 0.22 | 0.2 | [16 empty] | 21
| 1 | 3 | 0.4 | 0.35 | 0.29 | 0.19 | 0.22 | [16 empty] | 16
| 1 | 4 | 0.4 | 0.355 | 0.28 | 0.21 | 0.2 | [16 empty] | 0.29017
| 1 | 5 | 0.4 | 0.355 | 0.27 | 0.21 | 0.22 | [16 empty] | 27
| 1 | 6 | 0.4 | 0.335 | 0.28 | 0.22 | 0.21
| Mean |  | 0.4 | 0.34917 | 0.28667 | 0.205 | 0.21
| Median |  | 0.4 | 0.3525 | 0.285 | 0.21 | 0.21
| Std Dev. |  | 6.0809e-17 | 0.0080104 | 0.012111 | 0.016432 | 0.0089443
| Rel.Std.Dev. |  | 1.52e-14% | 2.294% | 4.225% | 8.015% | 4.259%
| PDM3 |  | 37.85% | 20.33% | -1.206% | -29.35% | -27.63%
| Z-Score (Absolute) |  | 1 | 0.55 | 0 | 0.72 | 0.67
| Analytical results for Te in OREAS 926b (Certified Value 0.07 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 17 | 18 | 20 | 21 | 22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.063 | 0.07 | 0.07 | 0.08 | 0.07 | 0.08 | 0.13 | 0.08 | 0.1 | 0.06 | 0.2 | 0.11 | 0.1 | 0.05 | < 0.1 | 0.08 | [5 empty] | 1
| 1 | 2 | 0.067 | 0.06 | 0.08 | 0.07 | 0.13 | 0.08 | 0.47 | 0.09 | 0.07 | 0.06 | 0.2 | 0.09 | 0.11 | 0.09 | < 0.1 | 0.09 | [5 empty] | 12
| 1 | 3 | 0.063 | 0.07 | 0.06 | 0.06 | 0.06 | 0.05 | 0.24 | 0.09 | 0.07 | 0.07 | 0.2 | 0.1 | 0.08 | 0.06 | < 0.1 | 0.07 | [5 empty] | 16
| 1 | 4 | 0.066 | 0.07 | 0.07 | 0.07 | 0.07 | 0.06 | 0.46 | 0.06 | 0.09 | 0.06 | 0.2 | 0.1 | 0.09 | 0.09 | < 0.1 | 0.06 | [5 empty] | 0.069985
| 1 | 5 | 0.064 | 0.07 | 0.05 | 0.07 | 0.07 | 0.06 | 0.02 | 0.09 | 0.08 | 0.07 | 0.2 | 0.11 | 0.09 | 0.05 | < 0.1 | 0.12 | [5 empty] | 172
| 1 | 6 | 0.062 | 0.08 | 0.07 | 0.07 | 0.07 | 0.05 | 0.02 | 0.07 | < 0.03 | 0.06 | 0.2 | 0.09 | 0.13 | 0.07 | < 0.1 | 0.07
| Mean |  | 0.064167 | 0.07 | 0.066667 | 0.07 | 0.078333 | 0.063333 | 0.22333 | 0.08 | 0.082 | 0.063333 | 0.2 | 0.1 | 0.1 | 0.068333 |  | 0.081667
| Median |  | 0.0635 | 0.07 | 0.07 | 0.07 | 0.07 | 0.06 | 0.185 | 0.085 | 0.08 | 0.06 | 0.2 | 0.1 | 0.095 | 0.065 |  | 0.075
| Std Dev. |  | 0.0019408 | 0.0063246 | 0.010328 | 0.0063246 | 0.025626 | 0.013663 | 0.20422 | 0.012649 | 0.013038 | 0.005164 | 3.0405e-17 | 0.0089443 | 0.017889 | 0.018348 |  | 0.02137
| Rel.Std.Dev. |  | 3.025% | 9.035% | 15.49% | 9.035% | 32.71% | 21.57% | 91.44% | 15.81% | 15.9% | 8.154% | 1.52e-14% | 8.944% | 17.89% | 26.85% |  | 26.17%
| PDM3 |  | -8.313% | 0.02165% | -4.741% | 0.02165% | 11.93% | -9.504% | 219.1% | 14.31% | 17.17% | -9.504% | 185.8% | 42.89% | 42.89% | -2.36% |  | 16.69%
| Z-Score (Absolute) |  | 0.59 | 0 | 0.34 | 0 | 0.84 | 0.67 | 15.51 | 1.01 | 0.08 | 0.67 | NA | 3.03 | 3.03 | 0.17 | 2.02 | 1.18
Results from laboratory 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 926b (Certified Value 8.03 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 6.79 | 8.1 | 8.6 | 7.7 | 8.2 | 9.4 | 7.7 | 8.1 | 8.86 | 8.9923 | 9.15 | 8.2 | 6.62 | 6.5 | 6.6 | 7.6224 | 5.4 |  |  |  |  | 1
| 1 | 2 | 6.93 | 8 | 9 | 7.7 | 8 | 10 | 7.7 | 8.4 | 8.62 | 9.2234 | 9.25 | 8.3 | 6.82 | 7.4 | 6.8 | 7.5533 | 5.26 |  |  |  |  | 5
| 1 | 3 | 7.03 | 8 | 8 | 7.3 | 7.3 | 8.8 | 7.8 | 8.4 | 8.61 | 9.1563 | 9.15 | 8.3 | 6.71 | 6.9 | 7 | 7.6395 | 5.67 |  |  |  |  | 16
| 1 | 4 | 6.98 | 7.9 | 8.4 | 7.6 | 8.2 | 9.7 | 8.1 | 8.2 | 8.67 | 8.9212 | 9.3 | 8.4 | 6.39 | 8 | 7.1 | 7.7546 | 5.33 |  |  |  |  | 8.0296
| 1 | 5 | 7.12 | 7.8 | 8.2 | 7.8 | 7.9 | 9.2 | 8 | 8.1 | 8.89 | 9.138 | 9.45 | 8.5 | 6.54 | 8.1 | 6.8 | 7.7581 | 5.48 |  |  |  |  | 173
| 1 | 6 | 7.18 | 8.1 | 7.8 | 7.7 | 7.7 | 10.3 | 7.4 | 8.3 | 8.58 | 8.7897 | 9.3 | 8.5 | 6.66 | 8.3 | 6.7 | 7.6915 | 5.37
| Mean |  | 7.005 | 7.9833 | 8.3333 | 7.6333 | 7.8833 | 9.5667 | 7.7833 | 8.25 | 8.705 | 9.0368 | 9.2667 | 8.3667 | 6.6233 | 7.5333 | 6.8333 | 7.6699 | 5.4183
| Median |  | 7.005 | 8 | 8.3 | 7.7 | 7.95 | 9.55 | 7.75 | 8.25 | 8.645 | 9.0652 | 9.275 | 8.35 | 6.64 | 7.7 | 6.8 | 7.6655 | 5.385
| Std Dev. |  | 0.13925 | 0.1169 | 0.43205 | 0.17512 | 0.34303 | 0.5465 | 0.24833 | 0.13784 | 0.13517 | 0.16477 | 0.11255 | 0.12111 | 0.1476 | 0.72296 | 0.18619 | 0.08025 | 0.14331
| Rel.Std.Dev. |  | 1.988% | 1.464% | 5.185% | 2.294% | 4.351% | 5.713% | 3.191% | 1.671% | 1.553% | 1.823% | 1.215% | 1.447% | 2.229% | 9.597% | 2.725% | 1.046% | 2.645%
| PDM3 |  | -12.76% | -0.576% | 3.783% | -4.935% | -1.821% | 19.14% | -3.067% | 2.745% | 8.412% | 12.54% | 15.41% | 4.198% | -17.51% | -6.18% | -14.9% | -4.479% | -32.52%
| Z-Score (Absolute) |  | 1.23 | 0.14 | 0.63 | 0.35 | 0 | 2.35 | 0.14 | 0.51 | 1.15 | 1.61 | 1.93 | 0.67 | 1.76 | 0.49 | 1.46 | 0.3 | 3.44
| Analytical results for Ti in OREAS 926b (Indicative Value 48.9 ppm) | [22 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.5g |  | 0.5g | [5 empty] | 1
| 1 | 1 | 40 | < 50 | < 50 | < 50 | < 50 | < 100 | < 50 | 30 | 48 | < 50 | 40 | < 50 | 100 | < 100 | 50 | 18 | [5 empty] | 1
| 1 | 2 | 40 | < 50 | < 50 | < 50 | < 50 | < 100 | < 50 | 30 | 51 | < 50 | 40 | < 50 | < 100 | < 100 | 50 | 18 | [5 empty] | 22
| 1 | 3 | 30 | < 50 | < 50 | < 50 | < 50 | < 100 | < 50 | 30 | 48 | < 50 | 50 | < 50 | < 100 | < 100 | 70 | 18 | [5 empty] | 16
| 1 | 4 | 30 | < 50 | < 50 | < 50 | < 50 | < 100 | < 50 | 30 | 51 | < 50 | 50 | < 50 | < 100 | < 100 | 70 | 18 | [5 empty] | 48.8549
| 1 | 5 | 30 | < 50 | < 50 | < 50 | < 50 | < 100 | < 50 | 30 | 53 | < 50 | 50 | < 50 | < 100 | < 100 | 70 | 18 | [5 empty] | 28
| 1 | 6 | 40 | < 50 | < 50 | < 50 | < 50 | < 100 | < 50 | 30 | 53 | < 50 | 40 | < 50 | < 100 | < 100 | 70 | 18
| Mean |  | 35 | [6 empty] | 30 | 50.6667 |  | 45 |  | 100 |  | 63.3333 | 18
| Median |  | 35 | [6 empty] | 30 | 51 |  | 45 |  | 100 |  | 70 | 18
| Std Dev. |  | 5.4772 | [6 empty] | 0 | 2.2509 |  | 5.4772 |  |  |  | 10.328 | 0
| Rel.Std.Dev. |  | 15.65% | [6 empty] | 0% | 4.443% |  | 12.17% |  |  |  | 16.31% | 0%
| PDM3 |  | -28.36% | [6 empty] | -38.59% | 3.708% |  | -7.891% |  | 104.7% |  | 29.64% | -63.16%
| Z-Score (Absolute) |  | 1.01 | 0.34 | 0.34 | 0.34 | 0.34 | 3.03 | 0.34 | 0.34 | 3.12 | 0.34 | 2.36 | 0.34 | 4.16 | 3.03 | 4.83 | 1.28
| Analytical results for Tl in OREAS 926b (Certified Value 0.094 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 21 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | [5 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.09 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.1 | 0.12 | 0.09 | 0.14 | 0.1 | < 0.5 | 0.09 | 0.08 | 0.1 | 0.18647 | [5 empty] | 1
| 1 | 2 | 0.09 | 0.1 | 0.11 | 0.09 | 0.11 | 0.08 | 0.07 | 0.1 | 0.08 | 0.14 | 0.1 | < 0.5 | 0.09 | 0.06 | 0.08 | 0.15263 | [5 empty] | #N/A
| 1 | 3 | 0.088 | 0.1 | 0.09 | 0.09 | 0.11 | 0.09 | 0.07 | 0.1 | 0.09 | 0.14 | 0.1 | < 0.5 | 0.08 | 0.05 | 0.15 | 0.14092 | [5 empty] | 16
| 1 | 4 | 0.085 | 0.1 | 0.09 | 0.09 | 0.1 | 0.09 | 0.06 | 0.1 | 0.09 | 0.14 | 0.1 | < 0.5 | 0.08 | 0.06 | 0.11 | 0.14163 | [5 empty] | 0.093617
| 1 | 5 | 0.092 | 0.1 | 0.09 | 0.09 | 0.1 | 0.1 | 0.07 | 0.1 | 0.08 | 0.14 | 0.1 | < 0.5 | 0.08 | 0.06 | 0.1 | 0.14415 | [5 empty] | 174
| 1 | 6 | 0.088 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.06 | 0.09 | 0.08 | 0.15 | < 0.1 | < 0.5 | 0.08 | 0.06 | 0.08 | 0.11618
| Mean |  | 0.088833 | 0.1 | 0.096667 | 0.091667 | 0.10167 | 0.093333 | 0.071667 | 0.10167 | 0.085 | 0.14167 | 0.1 |  | 0.083333 | 0.061667 | 0.10333 | 0.147
| Median |  | 0.089 | 0.1 | 0.09 | 0.09 | 0.1 | 0.095 | 0.07 | 0.1 | 0.085 | 0.14 | 0.1 |  | 0.08 | 0.06 | 0.1 | 0.14289
| Std Dev. |  | 0.0024014 | 1.5202e-17 | 0.010328 | 0.0040825 | 0.0075277 | 0.008165 | 0.01472 | 0.0098319 | 0.0054772 | 0.0040825 | 0 |  | 0.005164 | 0.0098319 | 0.02582 | 0.022863
| Rel.Std.Dev. |  | 2.703% | 1.52e-14% | 10.68% | 4.454% | 7.404% | 8.748% | 20.54% | 9.671% | 6.444% | 2.882% | 0% |  | 6.197% | 15.94% | 24.99% | 15.55%
| PDM3 |  | -5.109% | 6.819% | 3.258% | -2.083% | 8.599% | -0.3027% | -23.45% | 8.599% | -9.204% | 51.33% | 6.819% |  | -10.98% | -34.13% | 10.38% | 57.02%
| Z-Score (Absolute) |  | 0.67 | 0.29 | 0 | 0.43 | 0.43 | 0.29 | 2.15 | 0.43 | 1 | 3.87 | NA | 13.2 | 1.15 | 3.01 | 0.57 | 4.33
Results from laboratory 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 926b (Indicative Value 0.069 ppm) | [22 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 08 | 15 | 17 | [18 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | [18 empty] | 3
|  |  | 0.5g | 40g | 0.5g | [18 empty] | 3
| 1 | 1 | < 0.1 | 0.075 | 0.07 | [18 empty] | 1
| 1 | 2 | 0.1 | 0.075 | 0.06 | [18 empty] | 17
| 1 | 3 | < 0.1 | 0.07 | 0.07 | [18 empty] | 16
| 1 | 4 | < 0.1 | 0.075 | 0.08 | [18 empty] | 0.069444
| 1 | 5 | < 0.1 | 0.075 | 0.06 | [18 empty] | 29
| 1 | 6 | 0.1 | 0.07 | 0.07
| Mean |  | 0.1 | 0.073333 | 0.068333
| Median |  | 0.1 | 0.075 | 0.07
| Std Dev. |  | 0 | 0.002582 | 0.0075277
| Rel.Std.Dev. |  | 0% | 3.521% | 11.02%
| PDM3 |  | 44% | 5.6% | -1.6%
| Z-Score (Absolute) |  | 0.67 | 2.02 | 0
| Analytical results for U in OREAS 926b (Certified Value 1.92 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.795 | 1.96 | 1.83 | 1.92 | 1.85 | 2.2 | 1.9 | 2 | 2.41 | 2.02 | 1.6696 | 2.05 | 2 | 1.77 | 1.9 | 1.71 | 1.938 | 1.98 |  |  |  | 1
| 1 | 2 | 1.875 | 2 | 1.93 | 1.91 | 1.86 | 2.3 | 1.82 | 1.98 | 2.46 | 2.04 | 1.6062 | 2.15 | 2 | 1.79 | 2.12 | 1.73 | 1.8625 | 1.89 |  |  |  | #N/A
| 1 | 3 | 1.855 | 1.98 | 1.91 | 1.73 | 1.82 | 2.1 | 1.82 | 1.93 | 2.57 | 2 | 1.6718 | 2.05 | 2 | 1.82 | 1.85 | 1.72 | 1.9112 | 1.97 |  |  |  | 16
| 1 | 4 | 1.855 | 1.97 | 1.95 | 1.79 | 1.87 | 2.2 | 1.88 | 1.9 | 2.59 | 2 | 1.6833 | 2.1 | 2.1 | 1.81 | 2.09 | 1.74 | 1.9371 | 1.89 |  |  |  | 1.9227
| 1 | 5 | 1.895 | 1.94 | 1.9 | 1.8 | 1.97 | 2.1 | 1.88 | 1.92 | 2.56 | 2.06 | 1.6581 | 2.15 | 2.1 | 1.83 | 2.13 | 1.68 | 1.9334 | 1.96 |  |  |  | 175
| 1 | 6 | 1.925 | 2 | 1.83 | 1.81 | 1.9 | 2.3 | 1.8 | 1.99 | 2.43 | 2 | 1.6668 | 2 | 2.1 | 1.87 | 2.11 | 1.71 | 1.9386 | 1.94
| Mean |  | 1.8667 | 1.975 | 1.8917 | 1.8267 | 1.8783 | 2.2 | 1.85 | 1.9533 | 2.5033 | 2.02 | 1.6593 | 2.0833 | 2.05 | 1.815 | 2.0333 | 1.715 | 1.9201 | 1.9383
| Median |  | 1.865 | 1.975 | 1.905 | 1.805 | 1.865 | 2.2 | 1.85 | 1.955 | 2.51 | 2.01 | 1.6682 | 2.075 | 2.05 | 1.815 | 2.1 | 1.715 | 1.9352 | 1.95
| Std Dev. |  | 0.044008 | 0.023452 | 0.050761 | 0.073937 | 0.051929 | 0.089443 | 0.041473 | 0.041793 | 0.078909 | 0.025298 | 0.027259 | 0.060553 | 0.054772 | 0.034496 | 0.12437 | 0.020736 | 0.030072 | 0.039707
| Rel.Std.Dev. |  | 2.358% | 1.187% | 2.683% | 4.048% | 2.765% | 4.066% | 2.242% | 2.14% | 3.152% | 1.252% | 1.643% | 2.907% | 2.672% | 1.901% | 6.116% | 1.209% | 1.566% | 2.049%
| PDM3 |  | -2.917% | 2.718% | -1.617% | -4.997% | -2.31% | 14.42% | -3.784% | 1.591% | 30.2% | 5.058% | -13.7% | 8.352% | 6.618% | -5.604% | 5.751% | -10.8% | -0.1357% | 0.8105%
| Z-Score (Absolute) |  | 0.44 | 0.32 | 0.26 | 0.72 | 0.36 | 1.89 | 0.55 | 0.17 | 4 | 0.63 | 1.88 | 1.07 | 0.84 | 0.8 | 0.73 | 1.49 | 0.06 | 0.06
| Analytical results for V in OREAS 926b (Certified Value 14.1 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 13.5 | 13 | 15 | 13 | 13 | 18 | 12 | 14 | 17 | 17.482 | 20 | 14 | 14.54 | 8 | 16.26 | 8.2 | 15.3606 | 11 |  |  |  | 1
| 1 | 2 | 14 | 13 | 15 | 13 | 13 | 19 | 12 | 14 | 17 | 17.341 | 20 | 14 | 14.76 | 10 | 16.91 | 8.5 | 15.2778 | 10.1 |  |  |  | 6
| 1 | 3 | 13.5 | 13 | 14 | 13 | 13 | 18 | 11 | 13 | 16 | 17.393 | 20 | 14 | 14.36 | 9 | 16.34 | 8.5 | 15.7865 | 11.4 |  |  |  | 16
| 1 | 4 | 13.6 | 13 | 14 | 13 | 13 | 19 | 11 | 13 | 17 | 17.484 | 20 | 14 | 13.03 | 10 | 16.38 | 8.6 | 16.2995 | 10.8 |  |  |  | 14.0718
| 1 | 5 | 13.9 | 13 | 14 | 13 | 13 | 17 | 12 | 13 | 17 | 17.388 | 20 | 14 | 13 | 10 | 16.26 | 8 | 16.5048 | 11.4 |  |  |  | 176
| 1 | 6 | 14 | 13 | 13 | 13 | 13 | 19 | 11 | 14 | 17 | 17.435 | 20 | 14 | 12.66 | 10 | 16.27 | 8.2 | 15.8294 | 11.2
| Mean |  | 13.75 | 13 | 14.1667 | 13 | 13 | 18.3333 | 11.5 | 13.5 | 16.8333 | 17.4205 | 20 | 14 | 13.725 | 9.5 | 16.4033 | 8.3333 | 15.8431 | 10.9833
| Median |  | 13.75 | 13 | 14 | 13 | 13 | 18.5 | 11.5 | 13.5 | 17 | 17.414 | 20 | 14 | 13.695 | 10 | 16.305 | 8.35 | 15.8079 | 11.1
| Std Dev. |  | 0.2429 | 0 | 0.75277 | 0 | 0 | 0.8165 | 0.54772 | 0.54772 | 0.40825 | 0.056846 | 0 | 0 | 0.92537 | 0.83666 | 0.25303 | 0.23381 | 0.49026 | 0.4916
| Rel.Std.Dev. |  | 1.767% | 0% | 5.314% | 0% | 0% | 4.454% | 4.763% | 4.057% | 2.425% | 0.3263% | 0% | 0% | 6.742% | 8.807% | 1.543% | 2.806% | 3.095% | 4.476%
| PDM3 |  | -2.287% | -7.616% | 0.6744% | -7.616% | -7.616% | 30.28% | -18.28% | -4.063% | 19.62% | 23.8% | 42.13% | -0.5101% | -2.464% | -32.49% | 16.57% | -40.78% | 12.59% | -21.95%
| Z-Score (Absolute) |  | 0 | 0.23 | 0.13 | 0.23 | 0.23 | 1.43 | 0.69 | 0.07 | 0.96 | 1.14 | 1.94 | 0.08 | 0 | 1.32 | 0.83 | 1.68 | 0.65 | 0.86
| Analytical results for W in OREAS 926b (Certified Value 3.18 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 20 | 22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 2.73 | 3.61 | 2.88 | 3.29 | 2.95 | 2.1 | 3.07 | 3.1 | 2.7 | 3.5 | 3.6 | 2.38 | 3.6 | 3.63 | 3.9 | [6 empty] | 1
| 1 | 2 | 2.78 | 3.76 | 2.86 | 3.24 | 3.03 | 2 | 3.02 | 3.03 | 2.7 | 3.5 | 3.6 | 2.45 | 4 | 3.69 | 3.9 | [6 empty] | 7
| 1 | 3 | 2.7 | 3.69 | 3.12 | 3.01 | 2.95 | 1.7 | 3.03 | 3.18 | 2.8 | 3.5 | 3.6 | 2.5 | 3.5 | 3.66 | 4 | [6 empty] | 16
| 1 | 4 | 2.77 | 3.61 | 3.1 | 3.02 | 2.99 | 1.9 | 3.07 | 3.21 | 2.7 | 3.5 | 3.7 | 2.95 | 3.8 | 3.75 | 3.7 | [6 empty] | 3.1781
| 1 | 5 | 2.7 | 3.64 | 3.08 | 3.09 | 3.1 | 2 | 3.23 | 3.17 | 2.8 | 3.5 | 3.6 | 2.8 | 3.9 | 3.7 | 4 | [6 empty] | 177
| 1 | 6 | 2.86 | 3.7 | 2.97 | 3.05 | 3.08 | 2.7 | 2.88 | 3.08 | 2.8 | 3.5 | 3.7 | 2.73 | 3.9 | 3.66 | 3.8
| Mean |  | 2.7567 | 3.6683 | 3.0017 | 3.1167 | 3.0167 | 2.0667 | 3.05 | 3.1283 | 2.75 | 3.5 | 3.6333 | 2.635 | 3.7833 | 3.6817 | 3.8833
| Median |  | 2.75 | 3.665 | 3.025 | 3.07 | 3.01 | 2 | 3.05 | 3.135 | 2.75 | 3.5 | 3.6 | 2.615 | 3.85 | 3.675 | 3.9
| Std Dev. |  | 0.060882 | 0.059133 | 0.11462 | 0.11928 | 0.064395 | 0.33862 | 0.11261 | 0.069113 | 0.054772 | 0 | 0.05164 | 0.22492 | 0.19408 | 0.041673 | 0.1169
| Rel.Std.Dev. |  | 2.209% | 1.612% | 3.818% | 3.827% | 2.135% | 16.39% | 3.692% | 2.209% | 1.992% | 0% | 1.421% | 8.536% | 5.13% | 1.132% | 3.01%
| PDM3 |  | -13.26% | 15.42% | -5.552% | -1.933% | -5.08% | -34.97% | -4.031% | -1.566% | -13.47% | 10.13% | 14.32% | -17.09% | 19.04% | 15.84% | 22.19%
| Z-Score (Absolute) |  | 0.63 | 0.97 | 0.2 | 0 | 0.18 | 1.85 | 0.12 | 0.02 | 0.64 | 0.67 | 0.91 | 0.85 | 1.17 | 0.99 | 1.35
| Analytical results for Y in OREAS 926b (Certified Value 5.04 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 4.94 | 5.18 | 5.64 | 5.02 | 5.02 | 5.52 | 4.41 | 5.56 | 6.0715 | 5.6 | 4.9 | 3.44 | 6.805 | 3.86 | 3.73 | [6 empty] | 1
| 1 | 2 | 5.13 | 5.15 | 5.7 | 4.83 | 5.12 | 6.11 | 4.22 | 5.7 | 6.0443 | 5.6 | 4.85 | 4.6 | 6.567 | 3.94 | 3.57 | [6 empty] | #N/A
| 1 | 3 | 4.97 | 4.92 | 5.36 | 4.84 | 5.07 | 5.09 | 4.45 | 5.67 | 5.7911 | 5.5 | 4.83 | 3.82 | 6.613 | 3.95 | 3.89 | [6 empty] | 16
| 1 | 4 | 5.03 | 5.06 | 5.32 | 4.81 | 5.04 | 5.43 | 4.38 | 5.77 | 5.7321 | 5.55 | 4.58 | 4.24 | 6.777 | 4 | 3.73 | [6 empty] | 5.0439
| 1 | 5 | 5.04 | 5 | 5.25 | 5.06 | 4.98 | 5.48 | 4.24 | 5.95 | 5.94 | 5.65 | 4.56 | 4.34 | 6.718 | 3.84 | 3.84 | [6 empty] | 178
| 1 | 6 | 5.24 | 5.21 | 5.03 | 5.01 | 4.91 | 5.87 | 4.1 | 5.67 | 5.793 | 5.3 | 4.33 | 4.26 | 6.674 | 3.86 | 3.76
| Mean |  | 5.0583 | 5.0867 | 5.3833 | 4.9283 | 5.0233 | 5.5833 | 4.3 | 5.72 | 5.8953 | 5.5333 | 4.675 | 4.1167 | 6.6923 | 3.9083 | 3.7533
| Median |  | 5.035 | 5.105 | 5.34 | 4.925 | 5.03 | 5.5 | 4.31 | 5.685 | 5.8665 | 5.575 | 4.705 | 4.25 | 6.696 | 3.9 | 3.745
| Std Dev. |  | 0.11053 | 0.11308 | 0.25033 | 0.11303 | 0.072847 | 0.35809 | 0.13491 | 0.13145 | 0.14369 | 0.12517 | 0.22206 | 0.41596 | 0.092584 | 0.064005 | 0.11039
| Rel.Std.Dev. |  | 2.185% | 2.223% | 4.65% | 2.294% | 1.45% | 6.414% | 3.137% | 2.298% | 2.437% | 2.262% | 4.75% | 10.1% | 1.383% | 1.638% | 2.941%
| PDM3 |  | 0.2864% | 0.8481% | 6.73% | -2.291% | -0.4075% | 10.7% | -14.75% | 13.4% | 16.88% | 9.704% | -7.314% | -18.38% | 32.68% | -22.51% | -25.59%
| Z-Score (Absolute) |  | 0 | 0.04 | 0.42 | 0.17 | 0.04 | 0.67 | 0.97 | 0.85 | 1.08 | 0.61 | 0.49 | 1.21 | 2.1 | 1.48 | 1.68
| Analytical results for Yb in OREAS 926b (Indicative Value 0.48 ppm) | [22 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 08 | 17 | 18 | 22 | [17 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 0.5g | 0.5g | 0.25g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.6 | 0.44 | 0.4 | 0.43 | [17 empty] | 1
| 1 | 2 | 0.6 | 0.46 | 0.5 | 0.4 | [17 empty] | 24
| 1 | 3 | 0.5 | 0.44 | 0.4 | 0.44 | [17 empty] | 16
| 1 | 4 | 0.6 | 0.49 | 0.5 | 0.42 | [17 empty] | 0.48042
| 1 | 5 | 0.5 | 0.45 | 0.5 | 0.45 | [17 empty] | 30
| 1 | 6 | 0.6 | 0.48 | 0.5 | 0.43
| Mean |  | 0.56667 | 0.46 | 0.46667 | 0.42833
| Median |  | 0.6 | 0.455 | 0.5 | 0.43
| Std Dev. |  | 0.05164 | 0.020976 | 0.05164 | 0.017224
| Rel.Std.Dev. |  | 9.113% | 4.56% | 11.07% | 4.021%
| PDM3 |  | 17.95% | -4.25% | -2.862% | -10.84%
| Z-Score (Absolute) |  | 3.64 | 0.12 | 0.12 | 1.23
| Analytical results for Zn in OREAS 926b (Certified Value 97 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 40g | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 118 | 94 | 91 | 86 | 90 | 101 | 404 | 93.2 | 171 | 95 | 519.0495 | 102 | 101 | 91 | 83 | 99.984 | 89.6 | 94.5097 | 87 |  |  | 1
| 1 | 2 | 97.2 | 90 | 90 | 89 | 86 | 108 | 161 | 97.6 | 102 | 98 | 184.9185 | 101 | 95 | 94 | 90 | 97.819 | 86.4 | 86.7463 | 85 |  |  | #N/A
| 1 | 3 | 92.3 | 90 | 95 | 99 | 87 | 88.9 | 128 | 91.8 | 97 | 94 | 144.336 | 102 | 93 | 95 | 73 | 90.39 | 91.1 | 84.5576 | 84 |  |  | 16
| 1 | 4 | 89.7 | 91 | 92 | 95 | 86 | 89.1 | 112 | 92 | 96 | 93 | 149.4915 | 98 | 95 | 93 | 84 | 91.957 | 90.3 | 87.6838 | 84 |  |  | 96.8167
| 1 | 5 | 88.2 | 90 | 88 | 93 | 89 | 94.5 | 114 | 93.6 | 93 | 94 | 124.6065 | 98 | 94 | 94 | 87 | 89.766 | 89.3 | 84.8142 | 84 |  |  | 179
| 1 | 6 | 89 | 91 | 86 | 90 | 89 | 88.3 | 107 | 90 | 98 | 93 | 92.4975 | 98 | 93 | 94 | 87 | 92.72 | 86 | 87.6343 | 82
| Mean |  | 95.7333 | 91 | 90.3333 | 92 | 87.8333 | 94.9667 | 171 | 93.0333 | 109.5 | 94.5 | 202.4832 | 99.8333 | 95.1667 | 93.5 | 84 | 93.7727 | 88.7833 | 87.6577 | 84.3333
| Median |  | 91 | 90.5 | 90.5 | 91.5 | 88 | 91.8 | 121 | 92.6 | 97.5 | 94 | 146.9137 | 99.5 | 94.5 | 94 | 85.5 | 92.3385 | 89.45 | 87.1903 | 84
| Std Dev. |  | 11.3862 | 1.5492 | 3.1411 | 4.6476 | 1.7224 | 8.0269 | 115.8136 | 2.5688 | 30.2704 | 1.8708 | 158.0339 | 2.0412 | 2.9944 | 1.3784 | 5.933 | 4.1678 | 2.099 | 3.6188 | 1.633
| Rel.Std.Dev. |  | 11.89% | 1.702% | 3.477% | 5.052% | 1.961% | 8.452% | 67.73% | 2.761% | 27.64% | 1.98% | 78.05% | 2.045% | 3.147% | 1.474% | 7.063% | 4.445% | 2.364% | 4.128% | 1.936%
| PDM3 |  | -1.119% | -6.008% | -6.697% | -4.975% | -9.279% | -1.911% | 76.62% | -3.908% | 13.1% | -2.393% | 109.1% | 3.116% | -1.704% | -3.426% | -13.24% | -3.144% | -8.297% | -9.46% | -12.89%
| Z-Score (Absolute) |  | 0.48 | 0.53 | 0.67 | 0.32 | 1.21 | 0.31 | 16.5 | 0.1 | 3.41 | 0.21 | 23.21 | 1.35 | 0.35 | 0 | 2.02 | 0.06 | 1 | 1.24 | 1.95
| Analytical results for Zr in OREAS 926b (Certified Value 8.79 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 13 | 15 | 17 | 18 | 19 | 22 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 40g | 0.5g | 0.25g | 0.25g | 0.5g | [8 empty] | 3
| 1 | 1 | 7.7 | 8.8 | 9 | 9 | 8.5 | 7.1 | 8.2 | 8.7 | 9 | 6.27 | 9.3 | 9.32 | 13.1 | [8 empty] | 1
| 1 | 2 | 7.97 | 8.7 | 9.1 | 8.8 | 8.5 | 7.2 | 8.9 | 9 | 10 | 6.57 | 7.9 | 10.22 | 13.1 | [8 empty] | 9
| 1 | 3 | 7.92 | 8.6 | 8.6 | 8.2 | 8.6 | 6.7 | 8.9 | 9 | 9 | 7.02 | 7.1 | 10.38 | 13.4 | [8 empty] | 16
| 1 | 4 | 7.76 | 8.7 | 9.1 | 8.5 | 8.5 | 6 | 9 | 9 | 9 | 6.8 | 8.5 | 10.61 | 12.8 | [8 empty] | 8.7863
| 1 | 5 | 7.84 | 8.8 | 8.7 | 8.8 | 8.2 | 6.6 | 9.3 | 9.1 | 9 | 6.47 | 7.9 | 10.29 | 13.5 | [8 empty] | 180
| 1 | 6 | 8 | 8.7 | 8.7 | 8.6 | 8.2 | 6.3 | 8.5 | 8.6 | 9 | 6.19 | 9.7 | 10.81 | 13.2
| Mean |  | 7.865 | 8.7167 | 8.8667 | 8.65 | 8.4167 | 6.65 | 8.8 | 8.9 | 9.1667 | 6.5533 | 8.4 | 10.2717 | 13.1833
| Median |  | 7.88 | 8.7 | 8.85 | 8.7 | 8.5 | 6.65 | 8.9 | 9 | 9 | 6.52 | 8.2 | 10.335 | 13.15
| Std Dev. |  | 0.11929 | 0.075277 | 0.22509 | 0.28107 | 0.17224 | 0.45935 | 0.38987 | 0.2 | 0.40825 | 0.31564 | 0.96954 | 0.51472 | 0.24833
| Rel.Std.Dev. |  | 1.517% | 0.8636% | 2.539% | 3.249% | 2.046% | 6.907% | 4.43% | 2.247% | 4.454% | 4.816% | 11.54% | 5.011% | 1.884%
| PDM3 |  | -10.49% | -0.7929% | 0.9143% | -1.552% | -4.207% | -24.31% | 0.1555% | 1.294% | 4.329% | -25.41% | -4.397% | 16.91% | 50.04%
| Z-Score (Absolute) |  | 2.01 | 0.08 | 0.45 | 0.08 | 0.65 | 4.99 | 0.29 | 0.53 | NA | 5.22 | 0.69 | 3.89 | 11.03
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.