## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 471
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate / Peroxide Fusion ICP
Al2O3, Aluminium(III) oxide (wt.%) | 16.3272 | 15.9738 | 16.6806 | 16.1103 | 16.5441
Ba, Barium (wt.%) | 0.24511 | 0.23733 | 0.25289 | 0.24111 | 0.24911
Be, Beryllium (ppm) | 1.9812 | 1.892 | 2.0705 | 1.8434 | 2.1191
Bi, Bismuth (ppm) | 0.30386 | 0.22908 | 0.37864 | IND | IND
CaO, Calcium oxide (wt.%) | 1.2312 | 1.188 | 1.2744 | 1.2082 | 1.2543
CeO2, Cerium(IV) oxide (ppm) | 251.2099 | 242.2264 | 260.1933 | 246.5763 | 255.8435
Co, Cobalt (ppm) | 15.3742 | 14.4078 | 16.3405 | 14.5843 | 16.164
Cr, Chromium (ppm) | 62.0951 | 56.4999 | 67.6903 | 59.2633 | 64.9269
Cs, Caesium (ppm) | 6.9678 | 6.5668 | 7.3689 | 6.6805 | 7.2552
Cu, Copper (ppm) | 35.6961 | 28.4782 | 42.9141 | 32.8226 | 38.5697
Dy2O3, Dysprosium(III) oxide (ppm) | 5.8112 | 5.4433 | 6.179 | 5.5677 | 6.0546
Er2O3, Erbium(III) oxide (ppm) | 2.7779 | 2.5893 | 2.9665 | 2.651 | 2.9048
Eu2O3, Europium(III) oxide (ppm) | 3.11 | 2.8886 | 3.3313 | 2.9806 | 3.2393
Fe2O3, Iron(III) oxide (wt.%) | 6.1723 | 6.0051 | 6.3395 | 6.0682 | 6.2764
Ga, Gallium (ppm) | 21.4698 | 20.1807 | 22.7589 | 20.4917 | 22.4479
Gd2O3, Gadolinium(III) oxide (ppm) | 9.4825 | 9.0297 | 9.9353 | 9.1405 | 9.8245
Hf, Hafnium (ppm) | 7.9281 | 7.3453 | 8.511 | 7.4704 | 8.3859
Ho2O3, Holmium(III) oxide (ppm) | 1.0281 | 0.96794 | 1.0883 | 0.98526 | 1.071
In, Indium (ppm) | < 0.1 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 3.2756 | 3.1869 | 3.3642 | 3.2167 | 3.3345
La2O3, Lanthanum(III) oxide (ppm) | 123.674 | 119.3581 | 127.9898 | 120.9378 | 126.4101
Li, Lithium (ppm) | 16.2893 | 13.5207 | 19.0579 | IND | IND
Lu2O3, Lutetium(III) oxide (ppm) | 0.36775 | 0.30775 | 0.42775 | 0.34689 | 0.38861
MgO, Magnesium oxide (wt.%) | 1.2998 | 1.2729 | 1.3268 | 1.2774 | 1.3223
MnO, Manganese oxide (wt.%) | 0.050498 | 0.048845 | 0.052151 | 0.049487 | 0.051509
Mo, Molybdenum (ppm) | 2.3297 | 1.9247 | 2.7346 | 2.2164 | 2.4429
Na2O, Sodium oxide (wt.%) | 1.2248 | 1.1779 | 1.2717 | 1.1976 | 1.2521
Nb, Niobium (ppm) | 12.6901 | 11.8759 | 13.5043 | 12.0404 | 13.3398
Nd2O3, Neodymium(III) oxide (ppm) | 96.3967 | 92.8855 | 99.9078 | 93.8855 | 98.9078
Ni, Nickel (ppm) | 31.8178 | 24.5543 | 39.0812 | 28.713 | 34.9225
P2O5, Phosphorus(V) oxide (wt.%) | 0.32172 | 0.30725 | 0.33619 | 0.31166 | 0.33179
Pb, Lead (ppm) | 33.0386 | 30.5387 | 35.5386 | 31.3258 | 34.7515
Pr6O11, Praseodymium(III,IV) oxide (ppm) | 27.6542 | 26.6018 | 28.7065 | 26.958 | 28.3503
Rb, Rubidium (ppm) | 117.4999 | 113.2063 | 121.7934 | 113.5635 | 121.4362
S, Sulphur (wt.%) | 0.036103 | 0.025635 | 0.04657 | IND | IND
Sb, Antimony (ppm) | 0.93333 | 0.75201 | 1.1147 | IND | IND
Sc, Scandium (ppm) | 13.3204 | 12.3815 | 14.2593 | 12.7415 | 13.8993
SiO2, Silicon dioxide (wt.%) | 63.3435 | 62.2152 | 64.4717 | 62.7176 | 63.9694
Sm2O3, Samarium(III) oxide (ppm) | 14.3531 | 13.565 | 15.1413 | 13.9365 | 14.7698
Sn, Tin (ppm) | 3.1671 | 2.4481 | 3.8862 | 2.8296 | 3.5047
Sr, Strontium (ppm) | 445.1241 | 431.0213 | 459.2269 | 435.4291 | 454.8191
Ta, Tantalum (ppm) | 0.88211 | 0.78241 | 0.98181 | IND | IND
Tb4O7, Terbium(III,IV) oxide (ppm) | 1.1851 | 1.0903 | 1.2798 | 1.1314 | 1.2388
Th, Thorium (ppm) | 22.2751 | 21.4649 | 23.0854 | 21.7631 | 22.7872
Ti, Titanium (wt.%) | 0.56583 | 0.54327 | 0.58838 | 0.55349 | 0.57817
Tl, Thallium (ppm) | 0.64193 | 0.54109 | 0.74278 | 0.60412 | 0.67975
Tm2O3, Thulium(III) oxide (ppm) | 0.39002 | 0.35742 | 0.42262 | 0.36795 | 0.41208
TREO, Total Rare Earth Oxides (ppm) | 537.2279 | 521.0205 | 553.4354 | 527.9334 | 546.5224
TREO+Y2O3, Total Rare Earth Oxides + Yttrium oxide (ppm) | 571.3211 | 555.9026 | 586.7397 | 561.6565 | 580.9858
U, Uranium (ppm) | 2.6086 | 2.4535 | 2.7637 | 2.5114 | 2.7058
V, Vanadium (ppm) | 103.9331 | 99.3448 | 108.5213 | 100.4171 | 107.449
W, Tungsten (ppm) | 4.2914 | 3.7696 | 4.8132 | 3.9463 | 4.6365
Y2O3, Yttrium(III) oxide (ppm) | 31.737 | 30.4768 | 32.9971 | 30.5797 | 32.8942
Yb2O3, Ytterbium(III) oxide (ppm) | 2.4554 | 2.3068 | 2.604 | 2.3611 | 2.5497
Zn, Zinc (ppm) | 78.3844 | 72.643 | 84.1258 | 73.899 | 82.8698
Zr, Zirconium (ppm) | 315.5706 | 298.2909 | 332.8503 | 304.2348 | 326.9065
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 16.4088 | 16.2064 | 16.6111 | 16.3479 | 16.4696
BaO, Barium oxide (wt.%) | 0.2722 | 0.2534 | 0.29099 | 0.26503 | 0.27936
CaO, Calcium oxide (wt.%) | 1.2056 | 1.1841 | 1.2271 | 1.1946 | 1.2166
CeO2, Cerium(IV) oxide (ppm) | 243.3942 | 136.789 | 349.9995 | 231.7832 | 255.0053
Dy2O3, Dysprosium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Eu2O3, Europium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 6.2031 | 6.1052 | 6.3011 | 6.1563 | 6.25
Gd2O3, Gadolinium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Ho2O3, Holmium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 3.2331 | 3.191 | 3.2752 | 3.2107 | 3.2555
La2O3, Lanthanum(III) oxide (ppm) | 145.4764 | 113.1508 | 177.8021 | 138.9696 | 151.9833
Lu2O3, Lutetium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
MgO, Magnesium oxide (wt.%) | 1.3012 | 1.2781 | 1.3243 | 1.286 | 1.3164
MnO, Manganese oxide (wt.%) | 0.048469 | 0.045152 | 0.051786 | 0.045735 | 0.051204
Na2O, Sodium oxide (wt.%) | 1.2238 | 1.1821 | 1.2655 | 1.2064 | 1.2412
P2O5, Phosphorus(V) oxide (wt.%) | 0.32433 | 0.3134 | 0.33525 | 0.318 | 0.33065
Pr6O11, Praseodymium(III,IV) oxide (ppm) | < 100 | IND | IND | IND | IND
SiO2, Silicon dioxide (wt.%) | 63.3514 | 62.6433 | 64.0596 | 63.1128 | 63.5901
Sm2O3, Samarium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
SO3, Sulphur trioxide (wt.%) | 0.067967 | 0.05712 | 0.078814 | IND | IND
Tb4O7, Terbium(III,IV) oxide (ppm) | < 100 | IND | IND | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.94247 | 0.91991 | 0.96502 | 0.93045 | 0.95449
Tm2O3, Thulium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
V2O5, Vanadium(V) oxide (ppm) | 174.597 | 133.9821 | 215.2118 | IND | IND
Yb2O3, Ytterbium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
ZnO, Zinc oxide (ppm) | 99.8418 | 73.7363 | 125.9473 | IND | IND
ZrO2, Zirconium dioxide (ppm) | 402.6519 | 305.4011 | 499.9027 | 386.3833 | 418.9205
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 4.7259 | 4.442 | 5.0098 | 4.662 | 4.7898
4-Acid Digestion
Al, Aluminium (wt.%) | 8.4609 | 8.1441 | 8.7777 | 8.2716 | 8.6502
As, Arsenic (ppm) | 5.2644 | 4.4302 | 6.0985 | 4.9099 | 5.6188
Ba, Barium (wt.%) | 0.24747 | 0.23819 | 0.25675 | 0.24343 | 0.25151
Be, Beryllium (ppm) | 1.9904 | 1.8845 | 2.0963 | 1.9061 | 2.0747
Bi, Bismuth (ppm) | 0.30067 | 0.28344 | 0.31789 | 0.28497 | 0.31637
Ca, Calcium (wt.%) | 0.86171 | 0.83192 | 0.89149 | 0.84536 | 0.87805
Ce, Cerium (ppm) | 205.9242 | 197.0694 | 214.779 | 201.4455 | 210.403
Co, Cobalt (ppm) | 15.5266 | 14.793 | 16.2602 | 15.0663 | 15.9869
Cr, Chromium (ppm) | 50.5618 | 47.3096 | 53.814 | 48.9376 | 52.186
Cs, Caesium (ppm) | 6.9862 | 6.7429 | 7.2295 | 6.8146 | 7.1578
Cu, Copper (ppm) | 32.7172 | 31.4984 | 33.9361 | 31.7587 | 33.6758
Dy, Dysprosium (ppm) | 4.8526 | 4.5973 | 5.1079 | 4.6867 | 5.0186
Er, Erbium (ppm) | 2.2995 | 2.1853 | 2.4137 | 2.2076 | 2.3914
Eu, Europium (ppm) | 2.6298 | 2.5033 | 2.7562 | 2.5373 | 2.7223
Fe, Iron (wt.%) | 4.2336 | 4.0772 | 4.39 | 4.1611 | 4.3061
Ga, Gallium (ppm) | 21.3324 | 20.4315 | 22.2332 | 20.8166 | 21.8481
Gd, Gadolinium (ppm) | 8.2033 | 7.7469 | 8.6597 | 7.9408 | 8.4658
Hf, Hafnium (ppm) | 4.2796 | 4.0288 | 4.5304 | 4.1196 | 4.4396
Ho, Holmium (ppm) | 0.84853 | 0.8015 | 0.89557 | 0.81689 | 0.88018
In, Indium (ppm) | 0.058768 | 0.051035 | 0.066502 | 0.054598 | 0.062939
K, Potassium (wt.%) | 2.6795 | 2.5891 | 2.7698 | 2.6162 | 2.7427
La, Lanthanum (ppm) | 103.0537 | 97.7902 | 108.3172 | 100.8234 | 105.284
Li, Lithium (ppm) | 16.3735 | 15.6383 | 17.1087 | 15.93 | 16.8169
Lu, Lutetium (ppm) | 0.3007 | 0.28085 | 0.32056 | 0.2865 | 0.31491
Mg, Magnesium (wt.%) | 0.75125 | 0.72001 | 0.78248 | 0.73641 | 0.76608
Mn, Manganese (wt.%) | 0.038195 | 0.037049 | 0.039341 | 0.037552 | 0.038838
Mo, Molybdenum (ppm) | 2.0702 | 1.9645 | 2.1759 | 1.9906 | 2.1499
Na, Sodium (wt.%) | 0.90258 | 0.86943 | 0.93573 | 0.88487 | 0.92029
Nb, Niobium (ppm) | 12.1126 | 11.5066 | 12.7186 | 11.6899 | 12.5353
Nd, Neodymium (ppm) | 81.8633 | 77.5166 | 86.2099 | 79.1952 | 84.5313
Ni, Nickel (ppm) | 32.6055 | 31.4098 | 33.8012 | 31.7826 | 33.4284
P, Phosphorus (wt.%) | 0.14307 | 0.13857 | 0.14757 | 0.13977 | 0.14637
Pb, Lead (ppm) | 32.123 | 30.6167 | 33.6292 | 31.1363 | 33.1096
Pr, Praseodymium (ppm) | 23.1426 | 21.9687 | 24.3165 | 22.4019 | 23.8832
Rb, Rubidium (ppm) | 119.3019 | 113.5173 | 125.0865 | 116.0797 | 122.5241
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.029965 | 0.028674 | 0.031256 | 0.028971 | 0.030959
Sb, Antimony (ppm) | 0.95703 | 0.90289 | 1.0112 | 0.92115 | 0.99291
Sc, Scandium (ppm) | 12.7025 | 12.1168 | 13.2881 | 12.3981 | 13.0069
Sm, Samarium (ppm) | 12.2566 | 11.7291 | 12.7841 | 11.824 | 12.6893
Sn, Tin (ppm) | 3.0466 | 2.8712 | 3.2221 | 2.9212 | 3.1721
Sr, Strontium (ppm) | 442.4787 | 429.3884 | 455.569 | 434.3871 | 450.5703
Ta, Tantalum (ppm) | 0.79103 | 0.72939 | 0.85267 | 0.74878 | 0.83327
Tb, Terbium (ppm) | 0.94623 | 0.88882 | 1.0036 | 0.90678 | 0.98567
Th, Thorium (ppm) | 23.1312 | 22.0539 | 24.2085 | 22.4568 | 23.8055
Ti, Titanium (wt.%) | 0.53068 | 0.51229 | 0.54908 | 0.51667 | 0.5447
Tl, Thallium (ppm) | 0.67882 | 0.63975 | 0.71789 | 0.6568 | 0.70084
Tm, Thulium (ppm) | 0.30278 | 0.2805 | 0.32505 | 0.28834 | 0.31721
U, Uranium (ppm) | 2.4423 | 2.3101 | 2.5745 | 2.3392 | 2.5454
V, Vanadium (ppm) | 101.9604 | 98.4106 | 105.5103 | 99.4639 | 104.457
W, Tungsten (ppm) | 3.9908 | 3.8009 | 4.1808 | 3.8454 | 4.1363
Y, Yttrium (ppm) | 22.5706 | 21.519 | 23.6223 | 21.923 | 23.2182
Yb, Ytterbium (ppm) | 1.9605 | 1.8296 | 2.0914 | 1.8703 | 2.0507
Zn, Zinc (ppm) | 76.5278 | 73.197 | 79.8586 | 74.5952 | 78.4604
Zr, Zirconium (ppm) | 168.047 | 160.3122 | 175.7819 | 163.7758 | 172.3182
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 471
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate / Peroxide Fusion ICP
Ag | ppm | 0.195 | Cd | ppm | < 10 | Se | ppm | < 20
As | ppm | 5.0667 | Ge | ppm | 1.0508 | Te | ppm | < 1
B | ppm | 87.4167 | Re | ppm | < 0.1
Borate Fusion XRF
As2O3 | ppm | 18.4882 | Hg | ppm | < 20 | Rh | ppm | < 1
Au | ppm | < 20 | I | ppm | < 10 | Ru | ppm | < 1
Bi2O3 | ppm | < 11 | In | ppm | < 10 | Sb2O3 | ppm | < 36
Cd | ppm | < 10 | Ir | ppm | < 10 | Sc2O3 | ppm | 16.4954
Cl | wt.% | 0.2375 | MoO3 | ppm | < 1.5 | Se | ppm | < 1
Co3O4 | ppm | < 50 | Nb2O5 | ppm | 78.8403 | SnO2 | ppm | < 100
Cr2O3 | ppm | 78.1672 | Nd2O3 | ppm | 84.2166 | SrO | ppm | 540.9681
Cs2O | ppm | < 11 | NiO | ppm | 95.4911 | Ta2O5 | ppm | < 100
CuO | ppm | 52.2216 | Os | ppm | < 10 | Tc | ppm | < 1
Er2O3 | ppm | 104.5107 | PbO | ppm | 56.039 | Te | ppm | 40.6667
F | ppm | < 5000 | Pd | ppb | < 1000 | Th | ppm | < 50
Fr | ppm | < 30 | Pm | ppm | 117.5 | Tl | ppm | < 50
Ga2O3 | ppm | 116.7214 | Pt | ppb | < 1000 | U3O8 | ppm | < 100
Ge | ppm | < 30 | Ra | ppm | < 20 | WO3 | ppm | < 100
H2O- | wt.% | 1.1023 | Rb2O | ppm | 118.7583 | Y | ppm | 62.0712
HfO2 | ppm | < 100 | Re | ppm | < 10
Infrared Combustion
C | wt.% | 0.041667 | S | wt.% | 0.03
4-Acid Digestion
Ag | ppm | 0.049 | Ge | ppm | 0.27442 | Se | ppm | 1.6859
Cd | ppm | 0.027797 | Hg | ppm | 0.0125 | Te | ppm | 0.085053
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 471
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate / Peroxide Fusion ICP
Al2O3, wt.% | 16.3272 | 0.40242 | 15.5224 | 17.1321 | 15.12 | 17.5345 | 2.465% | 4.929% | 7.394% | 15.5109 | 17.1436
Ba, wt.% | 0.24511 | 0.0075652 | 0.22998 | 0.26024 | 0.22241 | 0.26781 | 3.086% | 6.173% | 9.259% | 0.23285 | 0.25737
Be, ppm | 1.9812 | 0.044709 | 1.8918 | 2.0707 | 1.8471 | 2.1154 | 2.257% | 4.513% | 6.77% | 1.8822 | 2.0803
Bi, ppm | 0.30386 | 0.03738 | 0.2291 | 0.37862 | 0.19172 | 0.416 | 12.3% | 24.6% | 36.9% | 0.28867 | 0.31905
CaO, wt.% | 1.2312 | 0.048188 | 1.1349 | 1.3276 | 1.0867 | 1.3758 | 3.914% | 7.827% | 11.74% | 1.1697 | 1.2928
CeO2, ppm | 251.2099 | 6.439 | 238.3319 | 264.0878 | 231.893 | 270.5268 | 2.563% | 5.126% | 7.69% | 238.6494 | 263.7704
Co, ppm | 15.3742 | 0.69668 | 13.9808 | 16.7675 | 13.2841 | 17.4642 | 4.531% | 9.063% | 13.59% | 14.6054 | 16.1429
Cr, ppm | 62.0951 | 3.6563 | 54.7825 | 69.4077 | 51.1262 | 73.064 | 5.888% | 11.78% | 17.66% | 58.9903 | 65.1998
Cs, ppm | 6.9678 | 0.37388 | 6.2201 | 7.7156 | 5.8462 | 8.0895 | 5.366% | 10.73% | 16.1% | 6.6195 | 7.3162
Cu, ppm | 35.6961 | 8.3405 | 19.0151 | 52.3771 | 10.6746 | 60.7176 | 23.37% | 46.73% | 70.1% | 33.9113 | 37.4809
Dy2O3, ppm | 5.8112 | 0.254 | 5.3032 | 6.3192 | 5.0492 | 6.5732 | 4.371% | 8.742% | 13.11% | 5.5206 | 6.1017
Er2O3, ppm | 2.7779 | 0.14445 | 2.489 | 3.0668 | 2.3445 | 3.2112 | 5.2% | 10.4% | 15.6% | 2.639 | 2.9168
Eu2O3, ppm | 3.11 | 0.20188 | 2.7062 | 3.5137 | 2.5043 | 3.7156 | 6.491% | 12.98% | 19.47% | 2.9545 | 3.2655
Fe2O3, wt.% | 6.1723 | 0.1367 | 5.8989 | 6.4457 | 5.7622 | 6.5824 | 2.215% | 4.43% | 6.644% | 5.8637 | 6.4809
Ga, ppm | 21.4698 | 1.2807 | 18.9084 | 24.0311 | 17.6278 | 25.3118 | 5.965% | 11.93% | 17.89% | 20.3963 | 22.5433
Gd2O3, ppm | 9.4825 | 0.56199 | 8.3585 | 10.6065 | 7.7965 | 11.1685 | 5.927% | 11.85% | 17.78% | 9.0084 | 9.9566
Hf, ppm | 7.9281 | 0.52298 | 6.8822 | 8.9741 | 6.3592 | 9.4971 | 6.597% | 13.19% | 19.79% | 7.5317 | 8.3246
Ho2O3, ppm | 1.0281 | 0.040822 | 0.94648 | 1.1098 | 0.90566 | 1.1506 | 3.97% | 7.941% | 11.91% | 0.97672 | 1.0795
In, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 3.2756 | 0.084192 | 3.1072 | 3.444 | 3.023 | 3.5281 | 2.57% | 5.141% | 7.711% | 3.1118 | 3.4394
La2O3, ppm | 123.674 | 4.0149 | 115.6443 | 131.7037 | 111.6294 | 135.7185 | 3.246% | 6.493% | 9.739% | 117.4903 | 129.8577
Li, ppm | 16.2893 | 1.4728 | 13.3436 | 19.2349 | 11.8708 | 20.7077 | 9.042% | 18.08% | 27.12% | 15.4748 | 17.1038
Lu2O3, ppm | 0.36775 | 0.032044 | 0.30366 | 0.43184 | 0.27162 | 0.46388 | 8.714% | 17.43% | 26.14% | 0.34936 | 0.38614
MgO, wt.% | 1.2998 | 0.02817 | 1.2435 | 1.3562 | 1.2153 | 1.3843 | 2.167% | 4.334% | 6.502% | 1.2348 | 1.3648
MnO, wt.% | 0.050498 | 0.0014591 | 0.04758 | 0.053416 | 0.046121 | 0.054875 | 2.889% | 5.779% | 8.668% | 0.047973 | 0.053023
Mo, ppm | 2.3297 | 0.62385 | 1.082 | 3.5774 | 0.4581 | 4.2012 | 26.78% | 53.56% | 80.34% | 2.2132 | 2.4461
Na2O, wt.% | 1.2248 | 0.040087 | 1.1446 | 1.305 | 1.1046 | 1.3451 | 3.273% | 6.546% | 9.819% | 1.1636 | 1.2861
Nb, ppm | 12.6901 | 0.95283 | 10.7844 | 14.5957 | 9.8316 | 15.5486 | 7.508% | 15.02% | 22.53% | 12.0556 | 13.3246
Nd2O3, ppm | 96.3967 | 3.1707 | 90.0553 | 102.7381 | 86.8846 | 105.9088 | 3.289% | 6.578% | 9.868% | 91.5768 | 101.2165
Ni, ppm | 31.8178 | 3.6981 | 24.4216 | 39.2139 | 20.7235 | 42.912 | 11.62% | 23.25% | 34.87% | 30.2269 | 33.4086
P2O5, wt.% | 0.32172 | 0.0106 | 0.30052 | 0.34292 | 0.28992 | 0.35352 | 3.295% | 6.59% | 9.884% | 0.30564 | 0.33781
Pb, ppm | 33.0386 | 2.1498 | 28.739 | 37.3383 | 26.5891 | 39.4881 | 6.507% | 13.01% | 19.52% | 31.3867 | 34.6906
Pr6O11, ppm | 27.6542 | 0.73511 | 26.1839 | 29.1244 | 25.4488 | 29.8595 | 2.658% | 5.316% | 7.975% | 26.2715 | 29.0369
Rb, ppm | 117.4999 | 4.3992 | 108.7015 | 126.2983 | 104.3023 | 130.6975 | 3.744% | 7.488% | 11.23% | 111.6249 | 123.3749
S, wt.% | 0.036103 | 0.0096916 | 0.01672 | 0.055486 | 0.0070281 | 0.065177 | 26.84% | 53.69% | 80.53% | 0.034298 | 0.037908
Sb, ppm | 0.93333 | 0.14044 | 0.65245 | 1.2142 | 0.51201 | 1.3547 | 15.05% | 30.09% | 45.14% | 0.88667 | 0.98
Sc, ppm | 13.3204 | 1.0214 | 11.2777 | 15.3631 | 10.2563 | 16.3845 | 7.668% | 15.34% | 23% | 12.6544 | 13.9864
SiO2, wt.% | 63.3435 | 0.72649 | 61.8905 | 64.7964 | 61.164 | 65.5229 | 1.147% | 2.294% | 3.441% | 60.1763 | 66.5106
Sm2O3, ppm | 14.3531 | 0.76993 | 12.8133 | 15.893 | 12.0434 | 16.6629 | 5.364% | 10.73% | 16.09% | 13.6355 | 15.0708
Sn, ppm | 3.1671 | 0.47863 | 2.2099 | 4.1244 | 1.7313 | 4.603 | 15.11% | 30.22% | 45.34% | 3.0088 | 3.3255
Sr, ppm | 445.1241 | 12.5513 | 420.0215 | 470.2267 | 407.4702 | 482.778 | 2.82% | 5.639% | 8.459% | 422.8679 | 467.3803
Ta, ppm | 0.88211 | 0.096107 | 0.68989 | 1.0743 | 0.59378 | 1.1704 | 10.9% | 21.79% | 32.69% | 0.838 | 0.92621
Tb4O7, ppm | 1.1851 | 0.087442 | 1.0102 | 1.36 | 0.92276 | 1.4474 | 7.379% | 14.76% | 22.14% | 1.1258 | 1.2443
Th, ppm | 22.2751 | 0.74362 | 20.7879 | 23.7624 | 20.0443 | 24.506 | 3.338% | 6.677% | 10.01% | 21.1614 | 23.3889
Ti, wt.% | 0.56583 | 0.026484 | 0.51286 | 0.6188 | 0.48638 | 0.64528 | 4.68% | 9.361% | 14.04% | 0.53754 | 0.59412
Tl, ppm | 0.64193 | 0.09743 | 0.44707 | 0.83679 | 0.34964 | 0.93422 | 15.18% | 30.36% | 45.53% | 0.60984 | 0.67403
Tm2O3, ppm | 0.39002 | 0.032378 | 0.32526 | 0.45477 | 0.29288 | 0.48715 | 8.302% | 16.6% | 24.91% | 0.37052 | 0.40952
TREO, ppm | 537.2279 | 10.9308 | 515.3663 | 559.0895 | 504.4355 | 570.0203 | 2.035% | 4.069% | 6.104% | 510.3665 | 564.0893
TREO+Y2O3, ppm | 571.3211 | 14.6844 | 541.9524 | 600.6899 | 527.268 | 615.3743 | 2.57% | 5.14% | 7.711% | 542.7551 | 599.8872
U, ppm | 2.6086 | 0.1033 | 2.402 | 2.8152 | 2.2987 | 2.9185 | 3.96% | 7.92% | 11.88% | 2.4782 | 2.739
V, ppm | 103.9331 | 5.5059 | 92.9213 | 114.9448 | 87.4155 | 120.4507 | 5.298% | 10.6% | 15.89% | 98.7364 | 109.1297
W, ppm | 4.2914 | 0.32426 | 3.6429 | 4.9399 | 3.3186 | 5.2642 | 7.556% | 15.11% | 22.67% | 4.0768 | 4.506
Y2O3, ppm | 31.737 | 1.1663 | 29.4044 | 34.0696 | 28.2381 | 35.2358 | 3.675% | 7.35% | 11.02% | 30.1501 | 33.3238
Yb2O3, ppm | 2.4554 | 0.15848 | 2.1384 | 2.7723 | 1.98 | 2.9308 | 6.454% | 12.91% | 19.36% | 2.3326 | 2.5782
Zn, ppm | 78.3844 | 4.9718 | 68.4408 | 88.3279 | 63.4691 | 93.2997 | 6.343% | 12.69% | 19.03% | 74.4652 | 82.3036
Zr, ppm | 315.5706 | 20.171 | 275.2287 | 355.9125 | 255.0578 | 376.0835 | 6.392% | 12.78% | 19.18% | 299.7921 | 331.3492
Borate Fusion XRF
Al2O3, wt.% | 16.4088 | 0.24563 | 15.9175 | 16.9 | 15.6719 | 17.1456 | 1.497% | 2.994% | 4.491% | 15.5883 | 17.2292
BaO, wt.% | 0.2722 | 0.017216 | 0.23776 | 0.30663 | 0.22055 | 0.32384 | 6.325% | 12.65% | 18.97% | 0.25859 | 0.28581
CaO, wt.% | 1.2056 | 0.01871 | 1.1682 | 1.243 | 1.1494 | 1.2617 | 1.552% | 3.104% | 4.656% | 1.1453 | 1.2659
CeO2, ppm | 243.3942 | 63.3986 | 116.5971 | 370.1914 | 53.1985 | 433.59 | 26.05% | 52.1% | 78.14% | 231.2245 | 255.564
Dy2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Eu2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Fe2O3, wt.% | 6.2031 | 0.10661 | 5.9899 | 6.4164 | 5.8833 | 6.523 | 1.719% | 3.437% | 5.156% | 5.893 | 6.5133
Gd2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ho2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 3.2331 | 0.050689 | 3.1317 | 3.3345 | 3.081 | 3.3852 | 1.568% | 3.136% | 4.703% | 3.0714 | 3.3948
La2O3, ppm | 145.4764 | 44.4381 | 56.6003 | 234.3525 | 12.1623 | 278.7906 | 30.55% | 61.09% | 91.64% | 138.2026 | 152.7503
Lu2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
MgO, wt.% | 1.3012 | 0.020613 | 1.26 | 1.3424 | 1.2394 | 1.363 | 1.584% | 3.168% | 4.752% | 1.2361 | 1.3663
MnO, wt.% | 0.048469 | 0.0034306 | 0.041608 | 0.055331 | 0.038178 | 0.058761 | 7.078% | 14.16% | 21.23% | 0.046046 | 0.050893
Na2O, wt.% | 1.2238 | 0.043953 | 1.1359 | 1.3117 | 1.0919 | 1.3556 | 3.592% | 7.183% | 10.77% | 1.1626 | 1.285
P2O5, wt.% | 0.32433 | 0.011275 | 0.30178 | 0.34688 | 0.2905 | 0.35815 | 3.476% | 6.953% | 10.43% | 0.30811 | 0.34054
Pr6O11, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SiO2, wt.% | 63.3514 | 0.88918 | 61.5731 | 65.1298 | 60.6839 | 66.019 | 1.404% | 2.807% | 4.211% | 60.1838 | 66.519
Sm2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SO3, wt.% | 0.067967 | 0.0060472 | 0.055872 | 0.080061 | 0.049825 | 0.086108 | 8.897% | 17.79% | 26.69% | 0.064568 | 0.071365
Tb4O7, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
TiO2, wt.% | 0.94247 | 0.025782 | 0.8909 | 0.99403 | 0.86512 | 1.0198 | 2.736% | 5.471% | 8.207% | 0.89534 | 0.98959
Tm2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
V2O5, ppm | 174.597 | 33.3458 | 107.9053 | 241.2886 | 74.5595 | 274.6344 | 19.1% | 38.2% | 57.3% | 165.8671 | 183.3268
Yb2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
ZnO, ppm | 99.8418 | 22.681 | 54.4799 | 145.2037 | 31.7989 | 167.8847 | 22.72% | 45.43% | 68.15% | 94.8497 | 104.8339
ZrO2, ppm | 402.6519 | 42.7423 | 317.1673 | 488.1365 | 274.425 | 530.8788 | 10.62% | 21.23% | 31.85% | 382.5193 | 422.7845
Thermogravimetry
LOI1000, wt.% | 4.7259 | 0.53491 | 3.6561 | 5.7957 | 3.1212 | 6.3306 | 11.32% | 22.64% | 33.96% | 4.4896 | 4.9622
4-Acid Digestion
Al, wt.% | 8.4609 | 0.24062 | 7.9797 | 8.9421 | 7.739 | 9.1828 | 2.844% | 5.688% | 8.532% | 8.0379 | 8.8839
As, ppm | 5.2644 | 1.1524 | 2.9595 | 7.5692 | 1.8071 | 8.7216 | 21.89% | 43.78% | 65.67% | 5.0011 | 5.5276
Ba, wt.% | 0.24747 | 0.010978 | 0.22551 | 0.26943 | 0.21454 | 0.2804 | 4.436% | 8.872% | 13.31% | 0.2351 | 0.25984
Be, ppm | 1.9904 | 0.10653 | 1.7773 | 2.2035 | 1.6708 | 2.31 | 5.352% | 10.7% | 16.06% | 1.8909 | 2.0899
Bi, ppm | 0.30067 | 0.018572 | 0.26352 | 0.33781 | 0.24495 | 0.35638 | 6.177% | 12.35% | 18.53% | 0.28563 | 0.3157
Ca, wt.% | 0.86171 | 0.025181 | 0.81134 | 0.91207 | 0.78616 | 0.93725 | 2.922% | 5.844% | 8.767% | 0.81862 | 0.90479
Ce, ppm | 205.9242 | 12.5309 | 180.8625 | 230.986 | 168.3316 | 243.5168 | 6.085% | 12.17% | 18.26% | 195.628 | 216.2204
Co, ppm | 15.5266 | 0.72088 | 14.0848 | 16.9684 | 13.364 | 17.6892 | 4.643% | 9.286% | 13.93% | 14.7503 | 16.3029
Cr, ppm | 50.5618 | 4.2576 | 42.0466 | 59.077 | 37.789 | 63.3346 | 8.421% | 16.84% | 25.26% | 48.0337 | 53.0899
Cs, ppm | 6.9862 | 0.32319 | 6.3398 | 7.6326 | 6.0166 | 7.9558 | 4.626% | 9.252% | 13.88% | 6.6369 | 7.3355
Cu, ppm | 32.7172 | 1.2096 | 30.298 | 35.1365 | 29.0883 | 36.3461 | 3.697% | 7.394% | 11.09% | 31.0814 | 34.3531
Dy, ppm | 4.8526 | 0.31301 | 4.2266 | 5.4787 | 3.9136 | 5.7917 | 6.45% | 12.9% | 19.35% | 4.61 | 5.0953
Er, ppm | 2.2995 | 0.15295 | 1.9936 | 2.6054 | 1.8407 | 2.7584 | 6.651% | 13.3% | 19.95% | 2.1845 | 2.4145
Eu, ppm | 2.6298 | 0.17023 | 2.2893 | 2.9703 | 2.1191 | 3.1405 | 6.473% | 12.95% | 19.42% | 2.4983 | 2.7613
Fe, wt.% | 4.2336 | 0.11069 | 4.0122 | 4.4549 | 3.9015 | 4.5656 | 2.614% | 5.229% | 7.843% | 4.0219 | 4.4452
Ga, ppm | 21.3324 | 0.82431 | 19.6837 | 22.981 | 18.8594 | 23.8053 | 3.864% | 7.728% | 11.59% | 20.2657 | 22.399
Gd, ppm | 8.2033 | 0.69618 | 6.811 | 9.5957 | 6.1148 | 10.2919 | 8.487% | 16.97% | 25.46% | 7.7932 | 8.6135
Hf, ppm | 4.2796 | 0.28954 | 3.7005 | 4.8587 | 3.411 | 5.1482 | 6.766% | 13.53% | 20.3% | 4.0656 | 4.4936
Ho, ppm | 0.84853 | 0.053592 | 0.74135 | 0.95572 | 0.68776 | 1.0093 | 6.316% | 12.63% | 18.95% | 0.80611 | 0.89096
In, ppm | 0.058768 | 0.0045598 | 0.049649 | 0.067888 | 0.045089 | 0.072448 | 7.759% | 15.52% | 23.28% | 0.05583 | 0.061707
K, wt.% | 2.6795 | 0.074101 | 2.5313 | 2.8277 | 2.4572 | 2.9018 | 2.766% | 5.531% | 8.297% | 2.5455 | 2.8134
La, ppm | 103.0537 | 7.1324 | 88.7889 | 117.3185 | 81.6565 | 124.4509 | 6.921% | 13.84% | 20.76% | 97.901 | 108.2064
Li, ppm | 16.3735 | 0.62258 | 15.1283 | 17.6186 | 14.5057 | 18.2412 | 3.802% | 7.605% | 11.41% | 15.5548 | 17.1921
Lu, ppm | 0.3007 | 0.014443 | 0.27182 | 0.32959 | 0.25738 | 0.34403 | 4.803% | 9.606% | 14.41% | 0.28567 | 0.31574
Mg, wt.% | 0.75125 | 0.031816 | 0.68762 | 0.81488 | 0.6558 | 0.84669 | 4.235% | 8.47% | 12.71% | 0.71368 | 0.78881
Mn, wt.% | 0.038195 | 0.00075159 | 0.036692 | 0.039698 | 0.03594 | 0.04045 | 1.968% | 3.936% | 5.903% | 0.036285 | 0.040105
Mo, ppm | 2.0702 | 0.079324 | 1.9116 | 2.2289 | 1.8322 | 2.3082 | 3.832% | 7.663% | 11.5% | 1.9667 | 2.1737
Na, wt.% | 0.90258 | 0.025436 | 0.85171 | 0.95345 | 0.82627 | 0.97889 | 2.818% | 5.636% | 8.455% | 0.85745 | 0.94771
Nb, ppm | 12.1126 | 0.70293 | 10.7068 | 13.5185 | 10.0038 | 14.2214 | 5.803% | 11.61% | 17.41% | 11.507 | 12.7183
Nd, ppm | 81.8633 | 4.4987 | 72.8658 | 90.8607 | 68.3671 | 95.3594 | 5.495% | 10.99% | 16.49% | 77.7701 | 85.9564
Ni, ppm | 32.6055 | 1.0232 | 30.5591 | 34.6519 | 29.5359 | 35.6751 | 3.138% | 6.276% | 9.414% | 30.9752 | 34.2358
P, wt.% | 0.14307 | 0.0037862 | 0.1355 | 0.15064 | 0.13171 | 0.15443 | 2.646% | 5.293% | 7.939% | 0.13591 | 0.15022
Pb, ppm | 32.123 | 1.6539 | 28.8151 | 35.4308 | 27.1612 | 37.0847 | 5.149% | 10.3% | 15.45% | 30.5168 | 33.7291
Pr, ppm | 23.1426 | 1.5938 | 19.955 | 26.3302 | 18.3612 | 27.9239 | 6.887% | 13.77% | 20.66% | 21.9854 | 24.2997
Rb, ppm | 119.3019 | 6.8707 | 105.5605 | 133.0432 | 98.6899 | 139.9138 | 5.759% | 11.52% | 17.28% | 113.3368 | 125.267
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.029965 | 0.00084764 | 0.02827 | 0.03166 | 0.027422 | 0.032508 | 2.829% | 5.658% | 8.486% | 0.028467 | 0.031463
Sb, ppm | 0.95703 | 0.040543 | 0.87594 | 1.0381 | 0.8354 | 1.0787 | 4.236% | 8.473% | 12.71% | 0.90918 | 1.0049
Sc, ppm | 12.7025 | 0.74437 | 11.2137 | 14.1912 | 10.4694 | 14.9356 | 5.86% | 11.72% | 17.58% | 12.0673 | 13.3376
Sm, ppm | 12.2566 | 0.82162 | 10.6134 | 13.8999 | 9.7918 | 14.7215 | 6.704% | 13.41% | 20.11% | 11.6438 | 12.8695
Sn, ppm | 3.0466 | 0.16014 | 2.7264 | 3.3669 | 2.5662 | 3.527 | 5.256% | 10.51% | 15.77% | 2.8943 | 3.199
Sr, ppm | 442.4787 | 11.4586 | 419.5614 | 465.396 | 408.1028 | 476.8546 | 2.59% | 5.179% | 7.769% | 420.3548 | 464.6026
Ta, ppm | 0.79103 | 0.086631 | 0.61777 | 0.96429 | 0.53114 | 1.0509 | 10.95% | 21.9% | 32.86% | 0.75148 | 0.83058
Tb, ppm | 0.94623 | 0.078383 | 0.78946 | 1.103 | 0.71108 | 1.1814 | 8.284% | 16.57% | 24.85% | 0.89891 | 0.99354
Th, ppm | 23.1312 | 1.2744 | 20.5824 | 25.6799 | 19.308 | 26.9543 | 5.509% | 11.02% | 16.53% | 21.9746 | 24.2877
Ti, wt.% | 0.53068 | 0.01982 | 0.49104 | 0.57032 | 0.47122 | 0.59014 | 3.735% | 7.47% | 11.2% | 0.50415 | 0.55722
Tl, ppm | 0.67882 | 0.051715 | 0.57539 | 0.78225 | 0.52368 | 0.83397 | 7.618% | 15.24% | 22.86% | 0.64488 | 0.71276
Tm, ppm | 0.30278 | 0.028612 | 0.24555 | 0.36 | 0.21694 | 0.38861 | 9.45% | 18.9% | 28.35% | 0.28764 | 0.31791
U, ppm | 2.4423 | 0.15773 | 2.1268 | 2.7578 | 1.9691 | 2.9155 | 6.458% | 12.92% | 19.37% | 2.3202 | 2.5644
V, ppm | 101.9604 | 3.3262 | 95.308 | 108.6129 | 91.9818 | 111.9391 | 3.262% | 6.525% | 9.787% | 96.8624 | 107.0585
W, ppm | 3.9908 | 0.21992 | 3.551 | 4.4307 | 3.331 | 4.6506 | 5.511% | 11.02% | 16.53% | 3.7913 | 4.1904
Y, ppm | 22.5706 | 2.0405 | 18.4896 | 26.6517 | 16.449 | 28.6923 | 9.041% | 18.08% | 27.12% | 21.4421 | 23.6992
Yb, ppm | 1.9605 | 0.13238 | 1.6957 | 2.2252 | 1.5633 | 2.3576 | 6.753% | 13.51% | 20.26% | 1.8625 | 2.0585
Zn, ppm | 76.5278 | 2.4589 | 71.61 | 81.4456 | 69.151 | 83.9046 | 3.213% | 6.426% | 9.639% | 72.7014 | 80.3542
Zr, ppm | 168.047 | 9.4375 | 149.172 | 186.9221 | 139.7344 | 196.3596 | 5.616% | 11.23% | 16.85% | 159.6447 | 176.4494
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 471
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PPP*XRF | pressed powder pellet with X-ray fluorescence spectroscopy
Text Values:
N.A. | Not Applicable (Lab 06)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 471
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
AGAT Laboratories, Calgary, Alberta, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Genalysis, Adelaide, SA, Australia
Lucid Laboratories Private Limited, Hyderabad, Telangana, India
Ontario Geological Survey, Sudbury, Ontario, Canada
Performance Laboratories, Ruwa, Harare, Zimbabwe
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Indo Mineral Research, Bungursari, West Java, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan
UIS Analytical Services, Centurion, South Africa

## Fusion ICP
| Analytical results for Ag in OREAS 471 (Indicative Value 0.195 ppm) | [28 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 01 | 06 | 09 | 18 | 21 | 27 | [21 empty] | ppm
|  |  | BF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [21 empty] | 3
3
| 1 | 1 | < 0.5 | < 1 | < 5 | < 5 | 0.31 | 0.07 | [21 empty] | 1
| 1 | 2 | < 0.5 | < 1 | < 5 | < 5 | 0.315 | 0.09 | [21 empty] | 1
| 1 | 3 | < 0.5 | < 1 | < 5 | < 5 | 0.305 | 0.09 | [21 empty] | 16
| 1 | 4 | < 0.5 | < 1 | < 5 | < 5 | 0.316 | 0.1 | [21 empty] | 0.195
| 1 | 5 | < 0.5 | < 1 | < 5 | < 5 | 0.301 | 0.09 | [21 empty] | 7
| 1 | 6 | < 0.5 | < 1 | < 5 | < 5 | 0.283 | 0.07
| Mean | [5 empty] | 0.305 | 0.085
| Median | [5 empty] | 0.3075 | 0.09
| Std Dev. | [5 empty] | 0.012215 | 0.012247
| Rel.Std.Dev. | [5 empty] | 4.005% | 14.41%
| PDM3 | [5 empty] | 56.41% | -56.41%
| Z-Score (Absolute) |  | 0.44 | 0.28 | 6.02 | 6.02 | 0.28 | 0.91
| Analytical results for Al2O3 in OREAS 471 (Certified Value 16.33 wt.%) | [28 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 23 | 24 | 26 | 27 | [10 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 16.26 | 16.2235 | 16.6584 | 16.609 | 15.92 | 16.42 | 15.966 | 16.552 | 17.346 | 17.42 | 16.17 | 16.117 | 16.74 | 17.6725 | 15.828 | 16.5746 | 16.212 | [10 empty] | 1
| 1 | 2 | 16.33 | 16.0989 | 17.2271 | 16.476 | 16.01 | 16.287 | 15.626 | 16.495 | 17.251 | 17.71 | 16.76 | 16.155 | 16.81 | 16.8902 | 15.87 | 16.4238 | 16.382 | [10 empty] | #N/A
| 1 | 3 | 16.2 | 16.145 | 16.313 | 16.59 | 16.11 | 16.42 | 15.569 | 16.646 | 16.401 | 17.52 | 16.48 | 16.25 | 16.86 | 17.4741 | 15.8603 | 16.6802 | 16.25 | [10 empty] | 16
| 1 | 4 | 16.26 | 15.9902 | 15.2762 | 16.609 | 16.09 | 16.212 | 15.91 | 16.722 | 16.099 | 17.47 | 15.9401 | 15.947 | 16.74 | 17.3399 | 15.8568 | 16.4714 | 16.174 | [10 empty] | 16.3272
| 1 | 5 | 16.25 | 16.1195 | 15.8025 | 16.646 | 15.94 | 16.552 | 15.456 | 16.401 | 17.062 | 17.81 | 16.6341 | 16.287 | 16.85 | 17.5308 | 15.8365 | 16.8172 | 16.287 | [10 empty] | 7
| 1 | 6 | 16.14 | 16.0954 | 16.5926 | 16.514 | 16.05 | 16.25 | 15.853 | 16.911 | 16.835 | 17.54 | 16.55 | 15.966 | 16.73 | 18.6154 | 15.6912 | 17.0942 | 16.363
| Mean |  | 16.24 | 16.1121 | 16.3116 | 16.574 | 16.02 | 16.3568 | 15.73 | 16.6212 | 16.8323 | 17.5783 | 16.4224 | 16.1203 | 16.7883 | 17.5871 | 15.8238 | 16.6769 | 16.278
| Median |  | 16.255 | 16.1092 | 16.4528 | 16.5995 | 16.03 | 16.3535 | 15.7395 | 16.599 | 16.9485 | 17.53 | 16.515 | 16.136 | 16.775 | 17.5024 | 15.8467 | 16.6274 | 16.2685
| Std Dev. |  | 0.064187 | 0.075972 | 0.68829 | 0.064952 | 0.077974 | 0.1291 | 0.20738 | 0.18109 | 0.4931 | 0.15012 | 0.30813 | 0.14119 | 0.059133 | 0.57051 | 0.066825 | 0.24923 | 0.082557
| Rel.Std.Dev. |  | 0.3952% | 0.4715% | 4.22% | 0.3919% | 0.4867% | 0.7893% | 1.318% | 1.089% | 2.929% | 0.854% | 1.876% | 0.8758% | 0.3522% | 3.244% | 0.4223% | 1.494% | 0.5072%
| PDM3 |  | -0.5342% | -1.318% | -0.09538% | 1.511% | -1.882% | 0.1814% | -3.658% | 1.8% | 3.094% | 7.663% | 0.5828% | -1.267% | 2.824% | 7.717% | -3.083% | 2.142% | -0.3014%
| Z-Score (Absolute) |  | 0.3 | 0.62 | 0.12 | 0.55 | 0.86 | 0 | 1.6 | 0.67 | 1.21 | 3.11 | 0.17 | 0.6 | 1.1 | 3.14 | 1.36 | 0.82 | 0.2
| Analytical results for As in OREAS 471 (Indicative Value 5.07 ppm) | [28 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 09 | 10 | 18 | 21 | 27 | [17 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [17 empty] | 2
2
| 1 | 1 | 5 | 5 | < 30 | 36 | < 20 | < 20 | 5 | < 20 | 22.552 | 6.4 | [17 empty] | 1
| 1 | 2 | 5 | < 5 | < 30 | 31 | < 20 | < 20 | 5 | < 20 | 23.473 | 5.3 | [17 empty] | 2
| 1 | 3 | 5 | 5 | < 30 | 37 | < 20 | < 20 | 7 | < 20 | 24.121 | 6.1 | [17 empty] | 16
| 1 | 4 | 5 | < 5 | < 30 | 40 | < 20 | < 20 | < 3 | < 20 | 25.56 | 5.5 | [17 empty] | 5.0667
| 1 | 5 | 5 | 5 | < 30 | 37 | < 20 | < 20 | 7 | < 20 | 20.443 | 6.2 | [17 empty] | 8
| 1 | 6 | 5 | 5 | < 30 | 40 | < 20 | < 20 | 5 | < 20 | 25.305 | 6.6
| Mean |  | 5 | 5 |  | 36.8333 |  |  | 5.8 |  | 23.5757 | 6.0167
| Median |  | 5 | 5 |  | 37 |  |  | 5 |  | 23.797 | 6.15
| Std Dev. |  | 0 | 0 |  | 3.3116 |  |  | 1.0954 |  | 1.9025 | 0.51153
| Rel.Std.Dev. |  | 0% | 0% |  | 8.991% |  |  | 18.89% |  | 8.07% | 8.502%
| PDM3 |  | -1.316% | -1.316% |  | 627% |  |  | 14.47% |  | 365.3% | 18.75%
| Z-Score (Absolute) |  | 0.68 | 0.79 | 0.68 | 3.65 | 0 | 0 | 0.67 | 0 | 1.85 | 0.54
| Analytical results for B in OREAS 471 (Indicative Value 87 ppm) | [28 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 03 | 04 | 06 | 08 | 09 | 18 | [21 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | [21 empty] | 0
0
| 1 | 1 | 66 | 133 | < 1000 | < 50 | < 50 | < 50 | [21 empty] | 1
| 1 | 2 | 63 | 107 | < 1000 | < 50 | < 50 | < 50 | [21 empty] | 3
| 1 | 3 | 65 | 119 | < 1000 | < 50 | < 50 | < 50 | [21 empty] | 16
| 1 | 4 | 59 | 116 | < 1000 | < 50 | < 50 | < 50 | [21 empty] | 87.4167
| 1 | 5 | 60 | 103 | < 1000 | < 50 | < 50 | < 50 | [21 empty] | 9
| 1 | 6 | 61 | 97 | < 1000 | < 50 | < 50 | < 50
| Mean |  | 62.3333 | 112.5
| Median |  | 62 | 111.5
| Std Dev. |  | 2.8048 | 12.9267
| Rel.Std.Dev. |  | 4.5% | 11.49%
| PDM3 |  | -28.69% | 28.69%
| Z-Score (Absolute) |  | 0.67 | 2.49 | 16.48 | 0.67 | 0.67 | 0.67
| Analytical results for Ba in OREAS 471 (Certified Value 0.245 wt.%) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 26 | 27 |  |  |  | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.2392 | 0.23992 | 0.2407 | 0.241 | 0.2448 | 0.2562 | 0.258 | 0.2466 | 0.2405 | 0.24538 | 0.261 | 0.245 | 0.244 | 0.262 | 0.235 | 0.24182 | 0.2475 | 0.2522 | 0.245 | 0.23667 | 0.245 | 0.25696 | 0.24488 | 0.2395 |  |  |  | 1
| 1 | 2 | 0.2397 | 0.23711 | 0.2535 | 0.227 | 0.2466 | 0.2553 | 0.251 | 0.2515 | 0.2418 | 0.24391 | 0.259 | 0.234 | 0.244 | 0.248 | 0.237 | 0.25226 | 0.2491 | 0.255 | 0.2429 | 0.23889 | 0.239 | 0.25752 | 0.23942 | 0.2428 |  |  |  | #N/A
| 1 | 3 | 0.2375 | 0.2382 | 0.2411 | 0.244 | 0.2478 | 0.257 | 0.255 | 0.248 | 0.2413 | 0.24465 | 0.255 | 0.246 | 0.233 | 0.244 | 0.229 | 0.24206 | 0.2414 | 0.2469 | 0.251 | 0.24566 | 0.245 | 0.25655 | 0.23793 | 0.2404 |  |  |  | 16
| 1 | 4 | 0.2392 | 0.24003 | 0.2246 | 0.238 | 0.2467 | 0.2573 | 0.254 | 0.255 | 0.2214 | 0.2477 | 0.26 | 0.247 | 0.242 | 0.251 | 0.238 | 0.24139 | 0.2439 | 0.249 | 0.2435 | 0.2395 | 0.242 | 0.25775 | 0.25199 | 0.2388 |  |  |  | 0.24511
| 1 | 5 | 0.24 | 0.23567 | 0.2355 | 0.237 | 0.2461 | 0.2536 | 0.259 | 0.2497 | 0.2425 | 0.24775 | 0.259 | 0.25 | 0.236 | 0.243 | 0.233 | 0.25239 | 0.2454 | 0.252 | 0.2581 | 0.23955 | 0.245 | 0.25676 | 0.23977 | 0.2401 |  |  |  | 8
| 1 | 6 | 0.2393 | 0.23674 | 0.2455 | 0.23 | 0.2463 | 0.2542 | 0.251 | 0.2508 | 0.2375 | 0.24512 | 0.252 | 0.247 | 0.237 | 0.237 | 0.237 | 0.24565 | 0.2466 | 0.25 | 0.2466 | 0.23146 | 0.245 | 0.25781 | 0.24647 | 0.2425
| Mean |  | 0.23915 | 0.23794 | 0.24015 | 0.23617 | 0.24638 | 0.2556 | 0.25467 | 0.25027 | 0.2375 | 0.24575 | 0.25767 | 0.24483 | 0.23933 | 0.2475 | 0.23483 | 0.24593 | 0.24565 | 0.25085 | 0.24785 | 0.23862 | 0.2435 | 0.25723 | 0.24341 | 0.24068
| Median |  | 0.23925 | 0.23765 | 0.2409 | 0.2375 | 0.24645 | 0.25575 | 0.2545 | 0.25025 | 0.2409 | 0.24525 | 0.259 | 0.2465 | 0.2395 | 0.246 | 0.236 | 0.24386 | 0.246 | 0.251 | 0.2458 | 0.2392 | 0.245 | 0.25724 | 0.24233 | 0.24025
| Std Dev. |  | 0.00086891 | 0.0017693 | 0.0097106 | 0.0064936 | 0.00097451 | 0.0015007 | 0.0033862 | 0.0029364 | 0.0080764 | 0.0016082 | 0.0034448 | 0.0055648 | 0.0046332 | 0.0085499 | 0.0033714 | 0.0051844 | 0.0027355 | 0.0028325 | 0.0057982 | 0.0046113 | 0.00251 | 0.00053795 | 0.0053748 | 0.0016216
| Rel.Std.Dev. |  | 0.3633% | 0.7436% | 4.044% | 2.75% | 0.3955% | 0.5871% | 1.33% | 1.173% | 3.401% | 0.6544% | 1.337% | 2.273% | 1.936% | 3.454% | 1.436% | 2.108% | 1.114% | 1.129% | 2.339% | 1.932% | 1.031% | 0.2091% | 2.208% | 0.6738%
| PDM3 |  | -2.431% | -2.923% | -2.024% | -3.649% | 0.5196% | 4.28% | 3.899% | 2.104% | -3.105% | 0.2619% | 5.123% | -0.1128% | -2.357% | 0.9751% | -4.193% | 0.3342% | 0.2204% | 2.342% | 1.118% | -2.646% | -0.6568% | 4.943% | -0.6939% | -1.806%
| Z-Score (Absolute) |  | 0.77 | 0.92 | 0.64 | 1.14 | 0.14 | 1.31 | 1.19 | 0.63 | 0.98 | 0.06 | 1.57 | 0.05 | 0.74 | 0.28 | 1.31 | 0.09 | 0.05 | 0.71 | 0.33 | 0.83 | 0.22 | 1.51 | 0.23 | 0.57
| Analytical results for Be in OREAS 471 (Certified Value 1.98 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 19 | 20 | 21 | 24 | 27 | [12 empty] | ppm
|  |  | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 2 | 0.8 | < 20 | < 5 | < 5 | 2 | 2 | 2 | 1.9 | 1.9583 | 1 | 2 | 2.189 | 1.1403 | 1.95 | [12 empty] | 1
| 1 | 2 | 2 | < 0.2 | < 20 | < 5 | < 5 | 2 | 2 | 2 | 1.8 | 1.9965 | 2 | 2 | 2.362 | 1.1716 | 2 | [12 empty] | #N/A
| 1 | 3 | 2 | < 0.2 | < 20 | < 5 | < 5 | 2 | 2 | 2 | 1.9 | 1.8562 | 3 | 2 | 2.557 | 1.18 | 1.95 | [12 empty] | 16
| 1 | 4 | 2 | < 0.2 | < 20 | < 5 | < 5 | 2 | 2 | 2 | 2.2 | 1.8336 | 2 | 2 | 2.424 | 1.1846 | 1.94 | [12 empty] | 1.9812
| 1 | 5 | 2 | < 0.2 | < 20 | < 5 | < 5 | 2 | 2 | 2 | 2.1 | 1.9918 | 2 | 2 | 2.42 | 1.1506 | 1.95 | [12 empty] | 9
| 1 | 6 | 2 | 0.6 | < 20 | < 5 | < 5 | 2 | 2 | 2 | 1.9 | 1.9532 | 3 | 2 | 2.11 | 1.1259 | 2.02
| Mean |  | 2 | 0.7 |  |  |  | 2 | 2 | 2 | 1.9667 | 1.9316 | 2.1667 | 2 | 2.3437 | 1.1588 | 1.9683
| Median |  | 2 | 0.7 |  |  |  | 2 | 2 | 2 | 1.9 | 1.9558 | 2 | 2 | 2.391 | 1.1611 | 1.95
| Std Dev. |  | 0 | 0.14142 |  |  |  | 0 | 0 | 0 | 0.15055 | 0.06972 | 0.75277 | 0 | 0.16532 | 0.023547 | 0.033116
| Rel.Std.Dev. |  | 0% | 20.2% |  |  |  | 0% | 0% | 0% | 7.655% | 3.609% | 34.74% | 0% | 7.054% | 2.032% | 1.682%
| PDM3 |  | 0.9468% | -64.67% |  |  |  | 0.9468% | 0.9468% | 0.9468% | -0.7356% | -2.506% | 9.359% | 0.9468% | 18.29% | -41.51% | -0.6515%
| Z-Score (Absolute) |  | 0 | 16.76 | 78.87 | 4.93 | 4.93 | 0 | 0 | 0 | 0.33 | 0.67 | 1.64 | 0 | 3.39 | 8.29 | 0.31
| Analytical results for Bi in OREAS 471 (Certified Value 0.3 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 09 | 10 | 17 | 18 | 21 | 27 | [16 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES/MS | [16 empty] | 2
2
| 1 | 1 | < 0.4 | 0.3 | < 0.5 | < 5 | 0.3 | 0.3 | 1.1 | 0.3425 | 0.4 | 0.299 | 0.3 | [16 empty] | 1
| 1 | 2 | < 0.4 | 0.3 | < 0.5 | < 5 | 0.3 | 0.3 | 1 | 0.4265 | 0.4 | 0.306 | 0.26 | [16 empty] | #N/A
| 1 | 3 | < 0.4 | 0.3 | < 0.5 | < 5 | 0.3 | 0.2 | 0.9 | 0.3454 | 0.4 | 0.306 | 0.27 | [16 empty] | 16
| 1 | 4 | < 0.4 | 0.4 | < 0.5 | < 5 | 0.3 | 0.3 | 0.9 | 0.3508 | 0.7 | 0.303 | 0.3 | [16 empty] | 0.30386
| 1 | 5 | < 0.4 | 0.3 | < 0.5 | < 5 | 0.3 | 0.2 | 1.6 | 0.3473 | 0.6 | 0.31 | 0.28 | [16 empty] | 10
| 1 | 6 | < 0.4 | 0.3 | < 0.5 | < 5 | 0.3 | 0.3 | 1.3 | 0.3765 | 0.4 | 0.3 | 0.29
| Mean |  |  | 0.31667 |  |  | 0.3 | 0.26667 | 1.1333 | 0.36483 | 0.48333 | 0.304 | 0.28333
| Median |  |  | 0.3 |  |  | 0.3 | 0.3 | 1.05 | 0.34905 | 0.4 | 0.3045 | 0.285
| Std Dev. |  |  | 0.040825 |  |  | 0 | 0.05164 | 0.27325 | 0.032618 | 0.13292 | 0.0041473 | 0.01633
| Rel.Std.Dev. |  |  | 12.89% |  |  | 0% | 19.36% | 24.11% | 8.941% | 27.5% | 1.364% | 5.764%
| PDM3 |  |  | 4.214% |  |  | -1.271% | -12.24% | 273% | 20.07% | 59.06% | 0.04571% | -6.756%
| Z-Score (Absolute) |  | 1.3 | 0.16 | 0.67 | 27.42 | 0.05 | 0.47 | 10.36 | 0.76 | 2.24 | 0 | 0.26
| Analytical results for CaO in OREAS 471 (Certified Value 1.23 wt.%) | [28 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 23 | 26 | 27 | [12 empty] | wt.%
|  |  | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 1.24 | 1.2088 | 1.217 | 1.2109 | 1.105 | 1.26 | 1.26 | 1.12 | 1.21 | 1.1557 | 1.4 | 1.24 | 1.2341 | 1.3187 | 1.231 | [12 empty] | 1
| 1 | 2 | 1.25 | 1.2517 | 1.231 | 1.2673 | 1.091 | 1.26 | 1.26 | 1.26 | 1.23 | 1.2063 | 1.4 | 1.24 | 1.2117 | 1.2845 | 1.245 | [12 empty] | #N/A
| 1 | 3 | 1.24 | 1.1937 | 1.189 | 1.2173 | 1.105 | 1.26 | 1.26 | 1.12 | 1.22 | 1.1851 | 1.4 | 1.24 | 1.2299 | 1.3396 | 1.231 | [12 empty] | 16
| 1 | 4 | 1.24 | 1.143 | 1.217 | 1.2487 | 1.063 | 1.26 | 1.26 | 1.12 | 1.21 | 1.1867 | 1.26 | 1.23 | 1.2299 | 1.3108 | 1.217 | [12 empty] | 1.2312
| 1 | 5 | 1.25 | 1.1359 | 1.217 | 1.2793 | 1.105 | 1.26 | 1.26 | 1.12 | 1.24 | 1.2008 | 1.26 | 1.24 | 1.2299 | 1.3319 | 1.231 | [12 empty] | 11
| 1 | 6 | 1.25 | 1.1966 | 1.231 | 1.3126 | 1.063 | 1.26 | 1.26 | 1.12 | 1.22 | 1.2017 | 1.26 | 1.24 | 1.2495 | 1.3614 | 1.245
| Mean |  | 1.245 | 1.1883 | 1.217 | 1.256 | 1.0887 | 1.26 | 1.26 | 1.1433 | 1.2217 | 1.1894 | 1.33 | 1.2383 | 1.2308 | 1.3245 | 1.2333
| Median |  | 1.245 | 1.1951 | 1.217 | 1.258 | 1.098 | 1.26 | 1.26 | 1.12 | 1.22 | 1.1938 | 1.33 | 1.24 | 1.2299 | 1.3253 | 1.231
| Std Dev. |  | 0.0054772 | 0.043218 | 0.015336 | 0.038627 | 0.020607 | 0 | 0 | 0.057155 | 0.01169 | 0.018592 | 0.076681 | 0.0040825 | 0.012065 | 0.026344 | 0.010539
| Rel.Std.Dev. |  | 0.4399% | 3.637% | 1.26% | 3.075% | 1.893% | 0% | 0% | 4.999% | 0.9569% | 1.563% | 5.766% | 0.3297% | 0.9802% | 1.989% | 0.8545%
| PDM3 |  | 1.118% | -3.49% | -1.157% | 2.01% | -11.58% | 2.336% | 2.336% | -7.14% | -0.7776% | -3.4% | 8.021% | 0.576% | -0.03309% | 7.573% | 0.17%
| Z-Score (Absolute) |  | 0.3 | 1.19 | 0.43 | 0.59 | 3.8 | 0.67 | 0.67 | 2.39 | 0.31 | 1.16 | 2.51 | 0.12 | 0.07 | 2.39 | 0
| Analytical results for Cd in OREAS 471 (Indicative Value < 10 ppm) | [28 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 04 | 06 | 07 | 08 | 09 | 10 | 21 | 27 | [19 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [19 empty] | 1
1
| 1 | 1 | < 5 | < 10 | < 0.5 | < 10 | < 1 | < 0.2 | < 0.1 | 0.13 | [19 empty] | 1
| 1 | 2 | < 5 | < 10 | < 0.5 | < 10 | < 1 | < 0.2 | < 0.1 | 0.11 | [19 empty] | 8
| 1 | 3 | < 5 | < 10 | < 0.5 | < 10 | < 1 | < 0.2 | < 0.1 | 0.12 | [19 empty] | 16
| 1 | 4 | < 5 | < 10 | < 0.5 | < 10 | < 1 | < 0.2 | < 0.1 | 0.12 | [19 empty] | < 10
| 1 | 5 | < 5 | < 10 | < 0.5 | < 10 | < 1 | < 0.2 | < 0.1 | 0.11 | [19 empty] | 7
| 1 | 6 | < 5 | < 10 | < 0.5 | < 10 | < 1 | < 0.2 | < 0.1 | 0.12
| Mean | [8 empty] | 0.11833
| Median | [8 empty] | 0.12
| Std Dev. | [8 empty] | 0.0075277
| Rel.Std.Dev. | [8 empty] | 6.361%
PDM3
| Z-Score (Absolute) |  | 4.78 | 10.4 | 0.28 | 10.4 | 0.28 | 0.62 | 0.73 | 0.58
| Analytical results for CeO2 in OREAS 471 (Certified Value 251 ppm) | [28 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 0
0
| 1 | 1 | 246.9084 | 248.6636 | 261.6492 | 259.1924 | 252.0677 | 270.248 | 238.3096 | 247.2769 | 256.7356 | 251.6992 | 256.7356 | 242.609 | 262.8776 | 260.4208 | 244.4516 | 246.1624 | 265.9486 | 252.559 | 246.5399 | 244.6199 | 291.6901 | 250.5936 | 306.6477 | 255.57 | 257.0789 | 257.0058 | 239.54 | 1
| 1 | 2 | 246.9084 | 256.7369 | 276.39 | 246.9084 | 256.6128 | 265.3344 | 249.3652 | 253.6646 | 271.4764 | 251.085 | 254.2788 | 245.0658 | 254.2788 | 249.3652 | 245.68 | 254.8161 | 263.9832 | 263.0004 | 245.4343 | 244.5978 | 309.7845 | 241.9948 | 304.8565 | 255.22 | 252.2318 | 259.3152 | 245.63 | #N/A
| 1 | 3 | 251.822 | 245.9013 | 253.0504 | 256.7356 | 255.2615 | 270.248 | 245.68 | 253.6646 | 248.1368 | 252.1905 | 248.1368 | 245.0658 | 243.8374 | 245.0658 | 243.2232 | 242.8371 | 266.6856 | 250.8393 | 251.9448 | 247.8825 | 298.9992 | 250.5936 | 308.3637 | 255.77 | 253.9677 | 259.8435 | 253.83 | 16
| 1 | 4 | 251.822 | 248.1764 | 240.7664 | 249.3652 | 253.7874 | 264.106 | 243.2232 | 250.4708 | 235.8528 | 250.7164 | 253.0504 | 244.4516 | 257.964 | 251.822 | 249.3652 | 244.0255 | 258.3325 | 249.3652 | 245.5572 | 244.8312 | 296.6405 | 241.9948 | 307.4076 | 255.86 | 249.3781 | 257.0918 | 245.61 | 251.2099
| 1 | 5 | 253.0504 | 255.1084 | 255.5072 | 265.3344 | 254.7702 | 267.7912 | 244.4516 | 255.3844 | 284.9888 | 252.1905 | 254.2788 | 248.1368 | 251.822 | 243.2232 | 248.1368 | 255.747 | 263.1233 | 254.7702 | 260.7893 | 243.6261 | 308.9246 | 243.2232 | 307.8508 | 255.11 | 251.302 | 257.7552 | 256.13 | 12
| 1 | 6 | 246.9084 | 243.9412 | 261.6492 | 253.0504 | 253.7874 | 265.3344 | 240.7664 | 252.1905 | 256.7356 | 248.751 | 244.4516 | 245.68 | 248.1368 | 238.3096 | 250.5936 | 249.3379 | 255.2615 | 247.2769 | 247.6454 | 243.7735 | 293.067 | 250.5936 | 308.2662 | 255.07 | 251.9211 | 261.035 | 248.36
| Mean |  | 249.5699 | 249.7546 | 258.1687 | 255.0977 | 254.3812 | 267.177 | 243.6327 | 252.1086 | 258.9877 | 251.1054 | 251.822 | 245.1682 | 253.1528 | 248.0344 | 246.9084 | 248.821 | 262.2225 | 252.9685 | 249.6518 | 244.8885 | 299.851 | 246.4989 | 307.2321 | 255.4333 | 252.6466 | 258.6744 | 248.1833
| Median |  | 249.3652 | 248.42 | 258.5782 | 254.893 | 254.2788 | 266.5628 | 243.8374 | 252.9276 | 256.7356 | 251.3921 | 253.6646 | 245.0658 | 253.0504 | 247.2155 | 246.9084 | 247.7502 | 263.5532 | 251.6992 | 247.0927 | 244.6088 | 297.8198 | 246.9084 | 307.6292 | 255.395 | 252.0765 | 258.5352 | 246.995
| Std Dev. |  | 2.9499 | 5.0939 | 11.7674 | 6.7581 | 1.5468 | 2.6631 | 3.8586 | 2.8847 | 17.2924 | 1.2963 | 4.5963 | 1.7977 | 6.822 | 7.6951 | 2.9069 | 5.4781 | 4.5012 | 5.5491 | 5.9593 | 1.5467 | 7.8059 | 4.5079 | 1.3228 | 0.34529 | 2.6298 | 1.6422 | 6.0484
| Rel.Std.Dev. |  | 1.182% | 2.04% | 4.558% | 2.649% | 0.6081% | 0.9968% | 1.584% | 1.144% | 6.677% | 0.5163% | 1.825% | 0.7332% | 2.695% | 3.102% | 1.177% | 2.202% | 1.717% | 2.194% | 2.387% | 0.6316% | 2.603% | 1.829% | 0.4305% | 0.1352% | 1.041% | 0.6348% | 2.437%
| PDM3 |  | -0.6528% | -0.5793% | 2.77% | 1.548% | 1.262% | 6.356% | -3.016% | 0.3578% | 3.096% | -0.04157% | 0.2437% | -2.405% | 0.7734% | -1.264% | -1.712% | -0.9509% | 4.384% | 0.7001% | -0.6202% | -2.516% | 19.36% | -1.875% | 22.3% | 1.681% | 0.5719% | 2.971% | -1.205%
| Z-Score (Absolute) |  | 0.44 | 0.4 | 1.04 | 0.51 | 0.39 | 2.59 | 1.46 | 0 | 1.18 | 0.17 | 0.05 | 1.19 | 0.18 | 0.7 | 0.89 | 0.56 | 1.74 | 0.15 | 0.42 | 1.24 | 8.2 | 0.96 | 9.47 | 0.57 | 0.09 | 1.13 | 0.67
| Analytical results for Co in OREAS 471 (Certified Value 15.4 ppm) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 10 | 11 | 17 | 19 | 20 | 21 | 23 | 24 | 27 | [13 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*MS | PF*OES/MS | [13 empty] | 1
2
| 1 | 1 | 16 | 13.3 | 14.5 | 20.66 | 16 | 18.2 | 15.8 | 15.5365 | 14 | 15.3 | 14.721 | 12.3 | 15.5135 | 15.5 | [13 empty] | 1
| 1 | 2 | 16 | 12.8 | 14.8 | 19.02 | 16 | 18.3 | 15.3 | 15.7421 | 15.8 | 15.4 | 16.107 | 11.7 | 15.6244 | 15.3 | [13 empty] | #N/A
| 1 | 3 | 15 | 13.1 | 14.7 | 20.64 | 16 | 20.1 | 15.1 | 15.0925 | 15 | 15.4 | 16.264 | 16 | 15.5188 | 15.8 | [13 empty] | 16
| 1 | 4 | 16 | 13.3 | 14.3 | 17.86 | 16 | 17.3 | 14.9 | 14.6219 | 15.6 | 15.3 | 15.472 | 17.7 | 15.6214 | 15.2 | [13 empty] | 15.3742
| 1 | 5 | 16 | 12.8 | 15.3 | 20.29 | 16 | 18.3 | 15.6 | 15.4203 | 14.9 | 16 | 15.752 | 17.1 | 15.4996 | 15.2 | [13 empty] | 13
| 1 | 6 | 16 | 12.6 | 14.7 | 17.33 | 16 | 17.8 | 15 | 15.8812 | 14.4 | 15.8 | 14.749 | 13.3 | 15.5169 | 15.5
| Mean |  | 15.8333 | 12.9833 | 14.7167 | 19.3 | 16 | 18.3333 | 15.2833 | 15.3824 | 14.95 | 15.5333 | 15.5108 | 14.6833 | 15.5491 | 15.4167
| Median |  | 16 | 12.95 | 14.7 | 19.655 | 16 | 18.25 | 15.2 | 15.4784 | 14.95 | 15.4 | 15.612 | 14.65 | 15.5179 | 15.4
| Std Dev. |  | 0.40825 | 0.29269 | 0.33714 | 1.46 | 0 | 0.94798 | 0.35449 | 0.46169 | 0.68629 | 0.29439 | 0.6613 | 2.5756 | 0.057561 | 0.23166
| Rel.Std.Dev. |  | 2.578% | 2.254% | 2.291% | 7.565% | 0% | 5.171% | 2.319% | 3.001% | 4.591% | 1.895% | 4.263% | 17.54% | 0.3702% | 1.503%
| PDM3 |  | 2.987% | -15.55% | -4.277% | 25.54% | 4.071% | 19.25% | -0.5907% | 0.05376% | -2.759% | 1.035% | 0.889% | -4.493% | 1.138% | 0.2765%
| Z-Score (Absolute) |  | NA | 5.44 | 1.54 | 8.78 | NA | 6.6 | 0.26 | 0.04 | 1.01 | 0.3 | 0.25 | 1.61 | 0.34 | 0.04
Results from laboratories 1 and 8 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 471 (Certified Value 62 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 23 | 27 | [6 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | BF*OES | PF*OES/MS | [6 empty] | 0
1
| 1 | 1 | 60 | 62.8294 | 55 | 63 | 70 | 88 | 50 | < 500 | 80 | 62 | 64 | 68 | 64 | 67 | 61 | 58.61 | < 500 | 62 | 49.848 | < 70 | 49 | [6 empty] | 1
| 1 | 2 | 60 | 66.1582 | 60 | 60 | < 70 | 87 | 50 | < 500 | 70 | 61 | 62 | 63 | 61 | 64 | 60 | 59.603 | < 500 | 62 | 45.169 | < 70 | 46 | [6 empty] | #N/A
| 1 | 3 | 60 | 64.0454 | 69 | 62 | < 70 | 88 | 50 | < 500 | 80 | 60 | 61 | 69 | 60 | 62 | 60 | 56.9047 | < 500 | 62 | 76.424 | < 70 | 41 | [6 empty] | 16
| 1 | 4 | 60 | 65.4775 | 65 | 61 | 70 | 88 | 50 | < 500 | 85 | 60 | 70 | 67 | 64 | 65 | 61 | 58.5896 | < 500 | 62 | 68.399 | < 70 | 45 | [6 empty] | 62.0951
| 1 | 5 | 60 | 59.5538 | 55 | 62 | 70 | 90 | 50 | < 500 | 70 | 61 | 65 | 70 | 64 | 61 | 60 | 61.3372 | < 500 | 68 | 52.791 | < 70 | 53 | [6 empty] | 14
| 1 | 6 | 60 | 62.1393 | 64 | 61 | < 70 | 85 | 50 | < 500 | 75 | 61 | 62 | 69 | 61 | 60 | 62 | 61.4423 | < 500 | 62 | 58.491 | < 70 | 49
| Mean |  | 60 | 63.3672 | 61.3333 | 61.5 | 70 | 87.6667 | 50 |  | 76.6667 | 60.8333 | 64 | 67.6667 | 62.3333 | 63.1667 | 60.6667 | 59.4145 |  | 63 | 58.5203 |  | 47.1667
| Median |  | 60 | 63.4374 | 62 | 61.5 | 70 | 88 | 50 |  | 77.5 | 61 | 63 | 68.5 | 62.5 | 63 | 60.5 | 59.1065 |  | 62 | 55.641 |  | 47.5
| Std Dev. |  | 0 | 2.4101 | 5.6804 | 1.0488 | 0 | 1.633 | 0 |  | 6.0553 | 0.75277 | 3.2863 | 2.5033 | 1.8619 | 2.6394 | 0.8165 | 1.759 |  | 2.4495 | 11.8695 |  | 4.1191
| Rel.Std.Dev. |  | 0% | 3.803% | 9.261% | 1.705% | 0% | 1.863% | 0% |  | 7.898% | 1.237% | 5.135% | 3.7% | 2.987% | 4.179% | 1.346% | 2.961% |  | 3.888% | 20.28% |  | 8.733%
| PDM3 |  | -3.374% | 2.049% | -1.227% | -0.9584% | 12.73% | 41.18% | -19.48% |  | 23.47% | -2.032% | 3.068% | 8.973% | 0.3837% | 1.726% | -2.3% | -4.317% |  | 1.457% | -5.757% |  | -24.04%
| Z-Score (Absolute) |  | NA | 0.33 | 0.13 | 0.1 | 2.42 | 5.89 | 2.72 | 43 | 3.37 | 0.25 | 0.48 | 1.31 | 0.1 | 0.29 | 0.29 | 0.57 | 43 | 0.18 | 0.78 | 6.15 | 3.37
Results from laboratory 1 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 471 (Certified Value 6.97 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 26 | 27 | [6 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 6.7 | 6.9008 | 7 | 6.6 | 6.9 | 6.7 | 6.7 | 7.3 | 7.29 | 6.75 | 6.78 | 7.51 | 7.04 | 6.5226 | 7.2 | 7 | 7.3 | 7.18 | 4.6 | 7.8451 | 6.37 | [6 empty] | 1
| 1 | 2 | 6.7 | 6.9465 | 6.9 | 7 | 6.85 | 6.5 | 7.1 | 7 | 7.64 | 6.35 | 6.86 | 7.04 | 6.95 | 6.7381 | 7.2 | 7 | 7.2 | 7.057 | 4.3 | 7.7251 | 6.54 | [6 empty] | #N/A
| 1 | 3 | 6.5 | 7.2542 | 7.1 | 6.7 | 7 | 6.7 | 6.3 | 7.2 | 7.27 | 6.87 | 6.42 | 7.6 | 7.23 | 6.392 | 7.1 | 6.8 | 7.4 | 7.067 | 5.8 | 7.7544 | 6.57 | [6 empty] | 16
| 1 | 4 | 6.6 | 7.1312 | 7 | 6.5 | 6.9 | 6.9 | 6 | 7.3 | 7.55 | 6.91 | 6.73 | 7.33 | 7.33 | 6.3633 | 6.9 | 7.2 | 7.1 | 7.292 | 6.5 | 7.4509 | 6.5 | [6 empty] | 6.9678
| 1 | 5 | 6.6 | 7.6738 | 7.3 | 6.7 | 6.99 | 6.7 | 7.4 | 7.6 | 7.56 | 6.99 | 6.75 | 7.05 | 6.63 | 6.7087 | 7.1 | 6.6 | 7.8 | 7.182 | 6.4 | 7.7037 | 6.45 | [6 empty] | 15
| 1 | 6 | 6.5 | 7.1018 | 7 | 6.7 | 6.74 | 6.8 | 6.5 | 6.8 | 7.45 | 6.79 | 6.27 | 6.76 | 7.39 | 6.7554 | 6.8 | 6.6 | 7.3 | 7.11 | 5.1 | 7.7506 | 6.67
| Mean |  | 6.6 | 7.1681 | 7.05 | 6.7 | 6.8967 | 6.7167 | 6.6667 | 7.2 | 7.46 | 6.7767 | 6.635 | 7.215 | 7.095 | 6.58 | 7.05 | 6.8667 | 7.35 | 7.148 | 5.45 | 7.705 | 6.5167
| Median |  | 6.6 | 7.1165 | 7 | 6.7 | 6.9 | 6.7 | 6.6 | 7.25 | 7.5 | 6.83 | 6.74 | 7.19 | 7.135 | 6.6157 | 7.1 | 6.9 | 7.3 | 7.145 | 5.45 | 7.7378 | 6.52
| Std Dev. |  | 0.089443 | 0.27909 | 0.13784 | 0.16733 | 0.096056 | 0.13292 | 0.5164 | 0.27568 | 0.15205 | 0.2258 | 0.23382 | 0.32042 | 0.28311 | 0.17772 | 0.16432 | 0.24221 | 0.2429 | 0.088546 | 0.92682 | 0.13351 | 0.10309
| Rel.Std.Dev. |  | 1.355% | 3.893% | 1.955% | 2.497% | 1.393% | 1.979% | 7.746% | 3.829% | 2.038% | 3.332% | 3.524% | 4.441% | 3.99% | 2.701% | 2.331% | 3.527% | 3.305% | 1.239% | 17.01% | 1.733% | 1.582%
| PDM3 |  | -5.279% | 2.874% | 1.179% | -3.844% | -1.021% | -3.605% | -4.322% | 3.332% | 7.063% | -2.744% | -4.777% | 3.547% | 1.825% | -5.566% | 1.179% | -1.452% | 5.485% | 2.586% | -21.78% | 10.58% | -6.475%
| Z-Score (Absolute) |  | 0.76 | 0.7 | 0.4 | 0.51 | 0 | 0.46 | 0.59 | 0.78 | 1.45 | 0.31 | 0.67 | 0.82 | 0.51 | 0.82 | 0.4 | 0.08 | 1.17 | 0.65 | 3.73 | 2.08 | 0.98
| Analytical results for Cu in OREAS 471 (Certified Value 35.7 ppm) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 10 | 17 | 21 | 23 | 24 | 27 | [15 empty] | ppm
|  |  | BF*OES/MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | PF*OES/MS | [15 empty] | 1
1
| 1 | 1 | 30 | 19 | 38 | 30 | 47 | 30 | 16 | 42.8215 | 33.35 | 38 | 29.328 | 36.8 | [15 empty] | 1
| 1 | 2 | 30 | 18 | 35 | 33 | 49 | 30 | 22 | 36.2922 | 35.513 | 37 | 29.728 | 37.5 | [15 empty] | #N/A
| 1 | 3 | 30 | 15 | 39 | 30 | 49 | 30 | 16 | 38.3789 | 35.619 | 48 | 29.184 | 36.3 | [15 empty] | 16
| 1 | 4 | 30 | 33 | 40 | 29 | 48 | 20 | 22 | 35.2241 | 34.875 | 71 | 29.36 | 36.1 | [15 empty] | 35.6961
| 1 | 5 | 30 | < 10 | 38 | 30 | 49 | 30 | 16 | 36.1248 | 34.438 | 53 | 29.6 | 36.7 | [15 empty] | 16
| 1 | 6 | 30 | 19 | 35 | 31 | 50 | 30 | 16 | 37.8128 | 32.614 | 51 | 29.528 | 37.2
| Mean |  | 30 | 20.8 | 37.5 | 30.5 | 48.6667 | 28.3333 | 18 | 37.7757 | 34.4015 | 49.6667 | 29.4547 | 36.7667
| Median |  | 30 | 19 | 38 | 30 | 49 | 30 | 16 | 37.0525 | 34.6565 | 49.5 | 29.444 | 36.75
| Std Dev. |  | 0 | 7.0143 | 2.0736 | 1.3784 | 1.0328 | 4.0825 | 3.0984 | 2.7299 | 1.2037 | 12.3882 | 0.19974 | 0.52789
| Rel.Std.Dev. |  | 0% | 33.72% | 5.53% | 4.519% | 2.122% | 14.41% | 17.21% | 7.227% | 3.499% | 24.94% | 0.6781% | 1.436%
| PDM3 |  | -15.96% | -41.73% | 5.053% | -14.56% | 36.34% | -20.63% | -49.57% | 5.826% | -3.627% | 39.14% | -17.48% | 2.999%
| Z-Score (Absolute) |  | NA | 2.09 | 0.23 | 0.61 | 1.57 | NA | 2.11 | 0.26 | 0.14 | 1.69 | 0.74 | 0.14
Results from laboratories 1 and 8 were removed due to their 10 ppm reading resolution.
| Analytical results for Dy2O3 in OREAS 471 (Certified Value 5.81 ppm) | [28 empty] | Certified
Dy2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 2
3
| 1 | 1 | 5.8533 | 6.2706 | 5.7385 | 6.0828 | 5.2565 | 5.7959 | 5.968 | 5.8533 | 6.3468 | 6.0828 | 5.6467 | 5.8188 | 5.6696 | 6.2664 | 5.75 | 5.6748 | 5.7385 | 5.4745 | 5.7844 | 5.3682 | 6.1228 | 6.0484 | 5.4671 | 5.34 | 5.3874 | 5.6811 | 5.38 | 1
| 1 | 2 | 5.968 | 6.1158 | 5.7385 | 5.6123 | 5.5549 | 5.8533 | 5.509 | 6.4271 | 6.5189 | 5.6237 | 5.8303 | 5.6696 | 6.3468 | 5.7959 | 5.6237 | 5.8387 | 5.7385 | 6.0484 | 5.8877 | 5.3074 | 5.6484 | 5.8533 | 5.3744 | 5.55 | 5.4514 | 5.5778 | 5.18 | #N/A
| 1 | 3 | 5.8533 | 6.2411 | 5.7385 | 5.9795 | 5.4745 | 5.8877 | 5.509 | 6.1976 | 6.0025 | 6.3123 | 5.727 | 5.9107 | 5.4401 | 5.6696 | 5.5549 | 5.6812 | 5.7385 | 5.5663 | 5.9451 | 5.5319 | 6.2506 | 5.968 | 5.3086 | 5.6 | 5.355 | 5.8074 | 5.66 | 16
| 1 | 4 | 5.7385 | 6.1891 | 4.5908 | 5.8418 | 5.6008 | 5.9336 | 5.8533 | 5.8533 | 6.0484 | 5.968 | 5.5319 | 5.5434 | 5.8418 | 5.9336 | 5.8533 | 5.4372 | 5.6237 | 5.4171 | 5.727 | 5.2864 | 6.5147 | 5.8992 | 5.2707 | 5.34 | 5.3164 | 5.8074 | 5.03 | 5.8112
| 1 | 5 | 5.968 | 6.7838 | 5.7385 | 6.1402 | 5.6811 | 5.7615 | 5.968 | 6.3123 | 7.1731 | 6.1976 | 5.8762 | 5.8188 | 5.6696 | 5.7729 | 6.1976 | 5.8631 | 5.7385 | 6.496 | 6.1058 | 5.3443 | 6.4069 | 5.7385 | 5.3196 | 5.37 | 5.6573 | 5.7385 | 5.18 | 17
| 1 | 6 | 5.8533 | 6.4498 | 5.7385 | 6.1287 | 5.7385 | 5.8647 | 5.509 | 6.0828 | 5.8647 | 5.8533 | 5.6008 | 5.5549 | 5.8992 | 5.4057 | 5.8074 | 5.7364 | 5.6237 | 6.0943 | 5.7959 | 5.4611 | 6.1054 | 6.0713 | 5.4012 | 5.25 | 5.4031 | 5.6582 | 5.29
| Mean |  | 5.8724 | 6.3417 | 5.5472 | 5.9642 | 5.551 | 5.8494 | 5.7194 | 6.1211 | 6.3257 | 6.0063 | 5.7022 | 5.7194 | 5.8112 | 5.8074 | 5.7978 | 5.7053 | 5.7002 | 5.8494 | 5.8743 | 5.3832 | 6.1748 | 5.9298 | 5.3569 | 5.4083 | 5.4284 | 5.7117 | 5.2867
| Median |  | 5.8533 | 6.2559 | 5.7385 | 6.0312 | 5.5778 | 5.859 | 5.6811 | 6.1402 | 6.1976 | 6.0254 | 5.6869 | 5.7442 | 5.7557 | 5.7844 | 5.7787 | 5.7088 | 5.7385 | 5.8074 | 5.8418 | 5.3563 | 6.1867 | 5.9336 | 5.347 | 5.355 | 5.3953 | 5.7098 | 5.235
| Std Dev. |  | 0.086396 | 0.24355 | 0.46855 | 0.20552 | 0.17157 | 0.062301 | 0.23427 | 0.23707 | 0.47954 | 0.24793 | 0.13393 | 0.15289 | 0.30767 | 0.28577 | 0.22565 | 0.153 | 0.059267 | 0.43015 | 0.13761 | 0.094906 | 0.30322 | 0.12565 | 0.071493 | 0.13615 | 0.12099 | 0.090273 | 0.21759
| Rel.Std.Dev. |  | 1.471% | 3.84% | 8.447% | 3.446% | 3.091% | 1.065% | 4.096% | 3.873% | 7.581% | 4.128% | 2.349% | 2.673% | 5.295% | 4.921% | 3.892% | 2.682% | 1.04% | 7.354% | 2.343% | 1.763% | 4.911% | 2.119% | 1.335% | 2.517% | 2.229% | 1.58% | 4.116%
| PDM3 |  | 1.054% | 9.13% | -4.542% | 2.634% | -4.476% | 0.6589% | -1.579% | 5.333% | 8.855% | 3.358% | -1.876% | -1.579% | 0.0005948% | -0.06524% | -0.2298% | -1.822% | -1.909% | 0.6589% | 1.087% | -7.364% | 6.258% | 2.041% | -7.816% | -6.932% | -6.586% | -1.711% | -9.026%
| Z-Score (Absolute) |  | 0.41 | 3.17 | 1.5 | 0.95 | 1.48 | 0.28 | 0.49 | 1.87 | 3.07 | 1.2 | 0.59 | 0.49 | 0.05 | 0.03 | 0.03 | 0.57 | 0.6 | 0.28 | 0.42 | 2.46 | 2.19 | 0.75 | 2.62 | 2.32 | 2.2 | 0.53 | NA
Results from Laboratory 1.28 were excluded due to a low bias observed in the Control sample.
| Analytical results for Er2O3 in OREAS 471 (Certified Value 2.78 ppm) | [28 empty] | Certified
Er2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 2.7444 | 3.042 | 2.3 | 2.9388 | 2.6186 | 2.813 | 2.7444 | 2.63 | 2.7787 | 2.63 | 2.8587 | 2.8359 | 2.9388 | 2.9388 | 2.63 | 2.7126 | 2.9731 | 2.5843 | 2.8816 | 2.7879 | 2.6513 | 2.9388 | 2.5757 | 3.0074 | 2.7155 | 2.6415 |  | 1
| 1 | 2 | 2.63 | 3.2471 | 2.3 | 2.7558 | 2.7444 | 2.8016 | 2.8587 | 2.9731 | 2.7215 | 2.9731 | 2.8587 | 2.8473 | 3.0417 | 2.6415 | 2.7215 | 2.7395 | 2.8587 | 3.0417 | 2.8473 | 2.7673 | 2.4457 | 2.8816 | 2.6801 | 3.0417 | 2.7064 | 2.7787 |  | 14
| 1 | 3 | 2.8587 | 3.076 | 3.4 | 2.7444 | 2.6987 | 2.7901 | 2.7444 | 2.8587 | 2.4928 | 2.8587 | 2.5729 | 2.8244 | 2.6186 | 2.8359 | 2.4814 | 2.6953 | 2.8587 | 2.6872 | 2.9731 | 2.8576 | 3.0942 | 2.8473 | 2.7614 | 3.0989 | 2.72 | 2.7101 |  | 16
| 1 | 4 | 2.7444 | 3.0389 | 2.3 | 2.8702 | 2.5729 | 2.5957 | 2.9731 | 2.9731 | 2.47 | 2.63 | 2.7787 | 2.813 | 2.6644 | 2.7673 | 2.7558 | 2.6541 | 2.8587 | 3.1103 | 2.9388 | 2.845 | 2.8015 | 2.7673 | 2.6114 | 3.2704 | 2.6387 | 2.7558 |  | 2.7779
| 1 | 5 | 2.7444 | 3.1712 | 2.3 | 2.8702 | 2.5386 | 2.7101 | 2.8587 | 2.63 | 3.2818 | 2.7444 | 2.8244 | 2.8359 | 2.8587 | 2.6529 | 2.5729 | 2.8019 | 2.8587 | 2.7673 | 3.0074 | 2.7879 | 3.1869 | 2.7787 | 2.6856 | 2.996 | 2.61 | 2.6529 |  | 18
| 1 | 6 | 2.7444 | 2.7825 | 2.3 | 2.9845 | 2.6872 | 2.7444 | 2.7444 | 2.7444 | 2.4814 | 2.7444 | 2.6529 | 2.9731 | 2.6872 | 2.5043 | 2.7558 | 2.7116 | 2.8587 | 3.1103 | 2.8587 | 2.7798 | 2.6983 | 2.8931 | 2.6672 | 2.9845 | 2.6919 | 2.7787
| Mean |  | 2.7444 | 3.0596 | 2.4833 | 2.8607 | 2.6434 | 2.7425 | 2.8206 | 2.8016 | 2.7044 | 2.7635 | 2.7577 | 2.8549 | 2.8016 | 2.7234 | 2.6529 | 2.7192 | 2.8778 | 2.8835 | 2.9178 | 2.8042 | 2.813 | 2.8511 | 2.6636 | 3.0665 | 2.6804 | 2.7196
| Median |  | 2.7444 | 3.059 | 2.3 | 2.8702 | 2.6529 | 2.7673 | 2.8016 | 2.8016 | 2.6072 | 2.7444 | 2.8016 | 2.8359 | 2.773 | 2.7101 | 2.6758 | 2.7121 | 2.8587 | 2.9045 | 2.9102 | 2.7879 | 2.7499 | 2.8645 | 2.6736 | 3.0246 | 2.6992 | 2.733
| Std Dev. |  | 0.072321 | 0.15837 | 0.44907 | 0.096059 | 0.07969 | 0.081636 | 0.093366 | 0.15762 | 0.31264 | 0.13368 | 0.11871 | 0.05905 | 0.17038 | 0.15535 | 0.11157 | 0.04933 | 0.046683 | 0.23215 | 0.065256 | 0.037444 | 0.28044 | 0.067327 | 0.064477 | 0.1081 | 0.045396 | 0.061544
| Rel.Std.Dev. |  | 2.635% | 5.176% | 18.08% | 3.358% | 3.015% | 2.977% | 3.31% | 5.626% | 11.56% | 4.837% | 4.305% | 2.068% | 6.081% | 5.704% | 4.206% | 1.814% | 1.622% | 8.051% | 2.236% | 1.335% | 9.97% | 2.361% | 2.421% | 3.525% | 1.694% | 2.263%
| PDM3 |  | -1.206% | 10.14% | -10.6% | 2.979% | -4.842% | -1.274% | 1.539% | 0.8526% | -2.646% | -0.5196% | -0.7254% | 2.774% | 0.8526% | -1.96% | -4.499% | -2.114% | 3.597% | 3.803% | 5.038% | 0.9486% | 1.263% | 2.636% | -4.116% | 10.39% | -3.509% | -2.098%
| Z-Score (Absolute) |  | 0.36 | 2.65 | 2.92 | 0.75 | 1.33 | 0.38 | 0.36 | 0.18 | 0.75 | 0.18 | 0.24 | 0.69 | 0.18 | 0.56 | 1.24 | 0.61 | 0.91 | 0.97 | 1.29 | 0.21 | 0.29 | 0.66 | 1.14 | 2.72 | 0.98 | 0.6
| Analytical results for Eu2O3 in OREAS 471 (Certified Value 3.11 ppm) | [28 empty] | Certified
Eu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 2.9295 | 3.3027 | 3.5 | 2.9642 | 3.2537 | 3.3695 | 3.2421 | 3.0105 | 3.0453 | 3.0105 | 2.7674 | 2.8832 | 3.1032 | 3.219 | 3.1148 | 2.9652 | 3.1263 | 3.6705 | 3.3695 | 3.2166 | 3.8524 | 3.1379 | 3.986 | 3.08 | 3.1479 | 2.8947 |  | 1
| 1 | 2 | 2.9642 | 3.0986 | 3.5 | 3.0105 | 3.1726 | 3.3695 | 3.2421 | 3.0105 | 3.1495 | 3.1263 | 2.7095 | 2.6747 | 2.779 | 2.9642 | 3.3 | 3.1782 | 3.0105 | 3.1611 | 3.3811 | 3.2722 | 3.7979 | 3.0221 | 4.0937 | 3.2884 | 3.12 | 3.0105 |  | #N/A
| 1 | 3 | 3.0105 | 3.1334 | 3.5 | 3.0569 | 3.2884 | 3.4621 | 3.1263 | 2.8947 | 2.9642 | 2.8947 | 2.7674 | 2.8716 | 3.0569 | 3.1032 | 2.9758 | 2.9692 | 3.0105 | 3.4853 | 3.4969 | 3.2954 | 4.1265 | 3.1032 | 4.0117 | 3.1263 | 3.0952 | 2.9758 |  | 16
| 1 | 4 | 3.0569 | 3.0706 | 2.3 | 3.1148 | 3.1032 | 3.3348 | 3.0105 | 2.8947 | 2.7558 | 3.1263 | 2.7326 | 2.7905 | 3.3695 | 3.0105 | 3.1379 | 2.9735 | 3.0105 | 3.5432 | 3.3463 | 3.2259 | 4.1481 | 3.0337 | 3.9542 | 3.0221 | 3.0605 | 2.9642 |  | 3.11
| 1 | 5 | 3.0337 | 3.2784 | 3.5 | 2.9411 | 3.3116 | 3.2769 | 3.0105 | 3.0105 | 3.3116 | 3.0105 | 2.8832 | 3.08 | 3.1032 | 2.8253 | 2.9874 | 3.0713 | 3.0105 | 3.4274 | 3.5316 | 3.2329 | 4.1627 | 3.1148 | 4.1168 | 3.08 | 3.1341 | 2.9642 |  | 19
| 1 | 6 | 2.999 | 3.1271 | 3.5 | 3.1032 | 3.1263 | 3.4042 | 3.2421 | 3.0105 | 2.9758 | 2.8947 | 2.7442 | 2.8716 | 3.2769 | 2.6863 | 2.7674 | 3.0115 | 2.8947 | 3.4737 | 3.3926 | 3.263 | 4.6486 | 3.0916 | 3.9914 | 3.2537 | 3.0708 | 2.9874
| Mean |  | 2.999 | 3.1685 | 3.3 | 3.0318 | 3.2093 | 3.3695 | 3.1456 | 2.9719 | 3.0337 | 3.0105 | 2.7674 | 2.8619 | 3.1148 | 2.9681 | 3.0472 | 3.0281 | 3.0105 | 3.4602 | 3.4197 | 3.251 | 4.1227 | 3.0839 | 4.0256 | 3.1418 | 3.1047 | 2.9662
| Median |  | 3.0048 | 3.1303 | 3.5 | 3.0337 | 3.2132 | 3.3695 | 3.1842 | 3.0105 | 3.0105 | 3.0105 | 2.7558 | 2.8716 | 3.1032 | 2.9874 | 3.0511 | 2.9925 | 3.0105 | 3.4795 | 3.3869 | 3.2479 | 4.1373 | 3.0974 | 4.0015 | 3.1032 | 3.1076 | 2.97
| Std Dev. |  | 0.046316 | 0.097457 | 0.4899 | 0.071908 | 0.087394 | 0.062569 | 0.11384 | 0.059794 | 0.18771 | 0.10357 | 0.060831 | 0.13291 | 0.20374 | 0.1912 | 0.18093 | 0.083774 | 0.073232 | 0.16882 | 0.075634 | 0.030667 | 0.30243 | 0.046123 | 0.064799 | 0.10604 | 0.035089 | 0.039038
| Rel.Std.Dev. |  | 1.544% | 3.076% | 14.85% | 2.372% | 2.723% | 1.857% | 3.619% | 2.012% | 6.187% | 3.44% | 2.198% | 4.644% | 6.541% | 6.442% | 5.938% | 2.767% | 2.433% | 4.879% | 2.212% | 0.9433% | 7.336% | 1.496% | 1.61% | 3.375% | 1.13% | 1.316%
| PDM3 |  | -3.569% | 1.881% | 6.111% | -2.514% | 3.195% | 8.345% | 1.147% | -4.438% | -2.452% | -3.197% | -11.02% | -7.975% | 0.1542% | -4.562% | -2.018% | -2.631% | -3.197% | 11.26% | 9.959% | 4.535% | 32.57% | -0.8387% | 29.44% | 1.023% | -0.168% | -4.624%
| Z-Score (Absolute) |  | 0.61 | 0.48 | 1.2 | 0.4 | 0.74 | 1.76 | 0.33 | 0.78 | 0.39 | 0.54 | 2.1 | 1.49 | 0.13 | 0.81 | 0.3 | 0.42 | 0.54 | 2.35 | 2.09 | 1 | 6.59 | 0.07 | 5.97 | 0.3 | 0.07 | 0.82
| Analytical results for Fe2O3 in OREAS 471 (Certified Value 6.17 wt.%) | [28 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 23 | 24 | 27 | [11 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS | [11 empty] | 2
3
| 1 | 1 | 6.2 | 6.2139 | 5.8929 | 6.119 | 6.3726 | 6.219 | 6.119 | 6.176 | 6.963 | 6.32 | 6.24 | 5.919 | 6.23 | 6.3007 | 5.7496 | 6.09 | [11 empty] | 1
| 1 | 2 | 6.22 | 6.1918 | 6.1094 | 6.119 | 6.3529 | 6.133 | 6.148 | 6.162 | 6.905 | 6.3 | 6.41 | 5.99 | 6.25 | 6.1291 | 5.7983 | 6.1 | [11 empty] | #N/A
| 1 | 3 | 6.15 | 6.0267 | 5.8047 | 6.076 | 6.189 | 6.205 | 6.219 | 6.119 | 6.648 | 6.29 | 6.24 | 6.019 | 6.16 | 6.2764 | 5.8078 | 6.53 | [11 empty] | 16
| 1 | 4 | 6.21 | 6.1548 | 5.6383 | 6.119 | 6.4169 | 5.976 | 6.076 | 6.248 | 6.434 | 6.25 | 6.19 | 5.933 | 6.17 | 6.1877 | 5.7906 | 5.98 | [11 empty] | 6.1723
| 1 | 5 | 6.2 | 6.0686 | 5.616 | 6.133 | 6.2811 | 6.248 | 6.219 | 6.119 | 6.863 | 6.38 | 6.45 | 6.048 | 6.21 | 6.2521 | 5.7919 | 6.26 | [11 empty] | 20
| 1 | 6 | 6.18 | 6.1332 | 5.9832 | 6.105 | 6.4708 | 5.919 | 6.105 | 6.305 | 6.834 | 6.29 | 6.37 | 5.948 | 6.2 | 6.2621 | 5.7659 | 6.89
| Mean |  | 6.1933 | 6.1315 | 5.8408 | 6.1118 | 6.3472 | 6.1167 | 6.1477 | 6.1882 | 6.7745 | 6.305 | 6.3167 | 5.9762 | 6.2033 | 6.2347 | 5.784 | 6.3083
| Median |  | 6.2 | 6.144 | 5.8488 | 6.119 | 6.3627 | 6.169 | 6.1335 | 6.169 | 6.8485 | 6.295 | 6.305 | 5.969 | 6.205 | 6.2571 | 5.7912 | 6.18
| Std Dev. |  | 0.025033 | 0.071962 | 0.19389 | 0.019661 | 0.10021 | 0.13759 | 0.059926 | 0.074338 | 0.19793 | 0.043243 | 0.1069 | 0.051207 | 0.034448 | 0.064045 | 0.021872 | 0.34313
| Rel.Std.Dev. |  | 0.4042% | 1.174% | 3.32% | 0.3217% | 1.579% | 2.249% | 0.9748% | 1.201% | 2.922% | 0.6859% | 1.692% | 0.8569% | 0.5553% | 1.027% | 0.3781% | 5.439%
| PDM3 |  | 0.3409% | -0.6611% | -5.371% | -0.9795% | 2.834% | -0.9012% | -0.3989% | 0.2572% | 9.757% | 2.15% | 2.339% | -3.177% | 0.503% | 1.011% | -6.291% | 2.204%
| Z-Score (Absolute) |  | 0.02 | 0.41 | 2.45 | 0.55 | 1.09 | 0.52 | 0.3 | 0.02 | 4.08 | 0.8 | 0.88 | 1.5 | 0.09 | 0.31 | 2.84 | 0.82
| Analytical results for Ga in OREAS 471 (Certified Value 21.5 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 26 | 27 |  |  |  |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*MS | PF*MS | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 22 | 22.5595 | 22.7 | 21.6 | 23 | 22 | 21 | 19 | 22.6 | 20.9 | 21.9 | 21.6 | 22.9 | 20.5 | 20.7147 | 22 | 16.5 | 19.8 | 20.307 | 16 | 23.1842 | 28.8389 | 19.74 |  |  |  |  | 1
| 1 | 2 | 22 | 23.8036 | 21.4 | 20.8 | 23 | 22 | 22 | 20 | 20.3 | 20.5 | 21.8 | 22.4 | 21.6 | 20.3 | 21.8838 | 22 | 16.4 | 20 | 20.679 | 15 | 22.7528 | 27.8371 | 19.2 |  |  |  |  | #N/A
| 1 | 3 | 22 | 22.1111 | 22.6 | 21.3 | 23 | 22 | 20 | 24 | 20.4 | 20.9 | 21.3 | 20.8 | 21.4 | 19.4 | 21.0075 | 23 | 15 | 20 | 20.988 | 21 | 22.6522 | 29.3574 | 20.06 |  |  |  |  | 16
| 1 | 4 | 22 | 22.3305 | 22.5 | 21.7 | 23 | 21 | 21 | 17 | 22.7 | 22.1 | 22.1 | 22.1 | 22.7 | 20.2 | 20.6685 | 22 | 14.2 | 20.1 | 21.2 | 22 | 22.3775 | 28.4213 | 20.15 |  |  |  |  | 21.4698
| 1 | 5 | 23 | 23.4836 | 22.1 | 20.9 | 24 | 22 | 22 | 20 | 24 | 21.6 | 20.8 | 20.7 | 22.3 | 20 | 21.5091 | 24 | 15.3 | 20.6 | 20.935 | 23 | 22.6625 | 28.0076 | 20.93 |  |  |  |  | 21
| 1 | 6 | 22 | 24.3794 | 21.6 | 21.6 | 21 | 23 | 21 | 19 | 22 | 20 | 21 | 21.2 | 20.9 | 20.6 | 20.9124 | 23 | 14.7 | 19.7 | 19.714 | 18 | 22.6746 | 28.7819 | 20.12
| Mean |  | 22.1667 | 23.1113 | 22.15 | 21.3167 | 22.8333 | 22 | 21.1667 | 19.8333 | 22 | 21 | 21.4833 | 21.4667 | 21.9667 | 20.1667 | 21.116 | 22.6667 | 15.35 | 20.0333 | 20.6372 | 19.1667 | 22.7173 | 28.5407 | 20.0333
| Median |  | 22 | 23.0215 | 22.3 | 21.45 | 23 | 22 | 21 | 19.5 | 22.3 | 20.9 | 21.55 | 21.4 | 21.95 | 20.25 | 20.9599 | 22.5 | 15.15 | 20 | 20.807 | 19.5 | 22.6685 | 28.6016 | 20.09
| Std Dev. |  | 0.40825 | 0.91001 | 0.54681 | 0.38687 | 0.98319 | 0.63246 | 0.75277 | 2.3166 | 1.4353 | 0.75366 | 0.52694 | 0.69186 | 0.78909 | 0.43205 | 0.48137 | 0.8165 | 0.92682 | 0.31411 | 0.54618 | 3.3116 | 0.26219 | 0.56703 | 0.5668
| Rel.Std.Dev. |  | 1.842% | 3.938% | 2.469% | 1.815% | 4.306% | 2.875% | 3.556% | 11.68% | 6.524% | 3.589% | 2.453% | 3.223% | 3.592% | 2.142% | 2.28% | 3.602% | 6.038% | 1.568% | 2.647% | 17.28% | 1.154% | 1.987% | 2.829%
| PDM3 |  | 3.246% | 7.646% | 3.168% | -0.7131% | 6.351% | 2.47% | -1.412% | -7.622% | 2.47% | -2.188% | 0.0632% | -0.01443% | 2.314% | -6.069% | -1.648% | 5.575% | -28.5% | -6.69% | -3.878% | -10.73% | 5.811% | 32.93% | -6.69%
| Z-Score (Absolute) |  | 0.57 | 1.34 | 0.56 | 0.12 | 1.11 | 0.43 | 0.24 | 1.33 | 0.43 | 0.38 | 0.01 | 0 | 0.41 | 1.06 | 0.29 | 0.98 | 4.97 | 1.17 | 0.67 | 1.87 | 1.02 | 5.75 | 1.17
| Analytical results for Gd2O3 in OREAS 471 (Certified Value 9.48 ppm) | [28 empty] | Certified
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 2
3
| 1 | 1 | 8.5292 | 9.2426 | 12.6786 | 10.0391 | 9.0825 | 9.9469 | 9.2208 | 9.3361 | 10.6039 | 9.2208 | 9.094 | 9.4167 | 9.209 | 9.8893 | 9.1286 | 9.2507 | 9.1055 | 8.5753 | 10.2927 | 8.7621 | 10.1373 | 9.1516 | 9.4498 | 9.4 | 10.457 | 9.9239 | 11.91 | 1
| 1 | 2 | 8.5292 | 9.5023 | 12.6786 | 9.5781 | 8.9211 | 9.6242 | 9.4513 | 9.7971 | 11.1572 | 9.5666 | 9.2208 | 9.2784 | 10.166 | 9.2208 | 8.9557 | 9.559 | 9.3361 | 8.7021 | 10.1775 | 8.7828 | 10.3493 | 9.1977 | 9.3649 | 9.67 | 10.173 | 10.2697 | 11.97 | #N/A
| 1 | 3 | 8.7598 | 9.8264 | 12.6786 | 9.3476 | 8.5753 | 9.9239 | 9.3361 | 9.3361 | 10.0046 | 9.6818 | 8.9672 | 9.3245 | 9.152 | 9.1516 | 8.4831 | 9.2274 | 9.5666 | 8.2296 | 10.5578 | 8.9903 | 10.3178 | 9.1055 | 9.5371 | 9.14 | 10.1683 | 10.1659 | 11.89 | 16
| 1 | 4 | 8.5292 | 9.4633 | 10.3734 | 9.8893 | 8.5523 | 10.1083 | 9.6818 | 9.6818 | 9.3476 | 9.7971 | 9.5205 | 9.0594 | 9.232 | 9.3822 | 9.509 | 9.107 | 9.2208 | 8.5869 | 10.2005 | 8.7517 | 11.1253 | 9.4398 | 9.4867 | 9.44 | 10.2034 | 9.9469 | 6.32 | 9.4825
| 1 | 5 | 8.7598 | 9.6067 | 11.526 | 9.97 | 8.7252 | 9.7856 | 9.5666 | 9.4513 | 11.6989 | 9.5666 | 9.1632 | 9.3822 | 9.751 | 9.0364 | 9.1171 | 9.6107 | 9.7971 | 8.4716 | 10.7768 | 8.7667 | 10.3465 | 9.2323 | 9.7224 | 9.29 | 10.2709 | 9.97 | 11.63 | 22
| 1 | 6 | 8.6445 | 9.4458 | 12.6786 | 10.3965 | 8.7713 | 10.1659 | 9.3361 | 8.9903 | 10.5693 | 9.3361 | 8.8289 | 9.1401 | 9.301 | 8.9903 | 9.532 | 9.3356 | 9.4513 | 8.633 | 10.189 | 8.8082 | 11.9176 | 9.5205 | 9.6691 | 9.17 | 10.2907 | 10.2466 | 10.96
| Mean |  | 8.6253 | 9.5145 | 12.1023 | 9.8701 | 8.7713 | 9.9258 | 9.4321 | 9.4321 | 10.5636 | 9.5282 | 9.1324 | 9.2669 | 9.4685 | 9.2784 | 9.1209 | 9.3484 | 9.4129 | 8.5331 | 10.3657 | 8.8103 | 10.699 | 9.2746 | 9.5383 | 9.3517 | 10.2605 | 10.0872 | 10.78
| Median |  | 8.5869 | 9.4828 | 12.6786 | 9.9296 | 8.7482 | 9.9354 | 9.3937 | 9.3937 | 10.5866 | 9.5666 | 9.1286 | 9.3015 | 9.2665 | 9.1862 | 9.1228 | 9.2931 | 9.3937 | 8.5811 | 10.2466 | 8.7747 | 10.3479 | 9.215 | 9.5119 | 9.345 | 10.2371 | 10.068 | 11.76
| Std Dev. |  | 0.11332 | 0.19351 | 0.96433 | 0.36722 | 0.20385 | 0.20102 | 0.16966 | 0.28622 | 0.82874 | 0.2146 | 0.23674 | 0.14022 | 0.40443 | 0.33013 | 0.38825 | 0.19788 | 0.24899 | 0.16682 | 0.24706 | 0.090352 | 0.68907 | 0.16682 | 0.13532 | 0.19651 | 0.10869 | 0.15811 | 2.2168
| Rel.Std.Dev. |  | 1.314% | 2.034% | 7.968% | 3.721% | 2.324% | 2.025% | 1.799% | 3.035% | 7.845% | 2.252% | 2.592% | 1.513% | 4.271% | 3.558% | 4.257% | 2.117% | 2.645% | 1.955% | 2.383% | 1.026% | 6.44% | 1.799% | 1.419% | 2.101% | 1.059% | 1.567% | 20.56%
| PDM3 |  | -9.04% | 0.3377% | 27.63% | 4.088% | -7.5% | 4.675% | -0.5313% | -0.5313% | 11.4% | 0.4816% | -3.692% | -2.274% | -0.1476% | -2.152% | -3.813% | -1.414% | -0.7339% | -10.01% | 9.314% | -7.089% | 12.83% | -2.193% | 0.5888% | -1.38% | 8.205% | 6.377% | 13.68%
| Z-Score (Absolute) |  | 1.75 | 0.18 | 5.79 | 0.95 | 1.43 | 1.07 | 0 | 0 | 2.45 | 0.21 | 0.65 | 0.36 | 0.08 | 0.33 | 0.67 | 0.18 | 0.04 | 1.95 | 2.02 | 1.35 | 2.75 | 0.34 | 0.23 | 0.17 | 1.8 | 1.42 | 2.92
| Analytical results for Ge in OREAS 471 (Indicative Value 1.05 ppm) | [28 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 04 | 06 | 08 | 10 | 17 | 21 | 27 | [19 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | PF*OES/MS | [19 empty] | 2
2
| 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1.1063 | 0.311 | 0.18 | [19 empty] | 1
| 1 | 2 | 1 | 2 | 1 | 1 | 2 | 1.2063 | 0.372 | 0.17 | [19 empty] | 18
| 1 | 3 | 1 | 2 | 1 | 1 | 2 | 1.1708 | 0.552 | 0.16 | [19 empty] | 16
| 1 | 4 | 1 | 2 | < 1 | 1 | 1 | 1.1036 | 0.48 | 0.18 | [19 empty] | 1.0508
| 1 | 5 | 1 | 2 | < 1 | 1 | 2 | 1.1195 | 0.384 | 0.16 | [19 empty] | 8
| 1 | 6 | 1 | 2 | 1 | 1 | 2 | 1.3093 | 0.304 | 0.17
| Mean |  | 1 | 2 | 1 | 1 | 1.6667 | 1.1693 | 0.4005 | 0.17
| Median |  | 1 | 2 | 1 | 1 | 2 | 1.1452 | 0.378 | 0.17
| Std Dev. |  | 0 | 0 | 0 | 0 | 0.5164 | 0.07967 | 0.097652 | 0.0089443
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 30.98% | 6.814% | 24.38% | 5.261%
| PDM3 |  | -4.835% | 90.33% | -4.835% | -4.835% | 58.61% | 11.28% | -61.89% | -83.82%
| Z-Score (Absolute) |  | 0 | 1.75 | 0.29 | 0 | 1.17 | 0.3 | 1.05 | 1.46
| Analytical results for Hf in OREAS 471 (Certified Value 7.93 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 06 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 27 | [9 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | BF*OES/MS | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 7.6 | 7.5265 | 7 | 8.66 | 7.7 | 8.5 | 8.42 | 7.81 | 7.78 | 8.62 | 8.12 | 7.9711 | 8.5 | 7.6 | 7.9 | 4.139 | 8 | 5.3 | [9 empty] | 1
| 1 | 2 | 7.5 | 7.3579 | 6.6 | 8.75 | 7.7 | 8.2 | 8.42 | 7.93 | 7.89 | 8.01 | 8.28 | 8.2331 | 7.9 | 7.9 | 7.8 | 4.295 | 8 | 5.2 | [9 empty] | #N/A
| 1 | 3 | 7.3 | 7.5612 | 7 | 8.54 | 8 | 8.5 | 8.1 | 7.95 | 7.08 | 8.46 | 7.8 | 7.9031 | 8.2 | 7.8 | 8.2 | 4.378 | 8 | 5.4 | [9 empty] | 16
| 1 | 4 | 7.3 | 8.0847 | 6.6 | 8.34 | 7.7 | 8.6 | 8.7 | 7.45 | 7.65 | 8.69 | 8 | 7.8582 | 8.4 | 8.5 | 7.9 | 4.248 | 8 | 5.6 | [9 empty] | 7.9281
| 1 | 5 | 7.3 | 7.6501 | 6.8 | 8.37 | 8 | 8.7 | 8.65 | 7.84 | 7.6 | 8.22 | 7.77 | 8.1652 | 8.3 | 7.8 | 8.3 | 4.541 | 8 | 5.4 | [9 empty] | 23
| 1 | 6 | 7.7 | 7.4813 | 6.6 | 9.28 | 8.1 | 8.5 | 8.13 | 7.55 | 7.19 | 7.71 | 8.38 | 8.1009 | 7.7 | 7.3 | 8 | 4.16 | 8 | 5.7
| Mean |  | 7.45 | 7.6103 | 6.7667 | 8.6567 | 7.8667 | 8.5 | 8.4033 | 7.755 | 7.5317 | 8.285 | 8.0583 | 8.0386 | 8.1667 | 7.8167 | 8.0167 | 4.2935 | 8 | 5.4333
| Median |  | 7.4 | 7.5438 | 6.7 | 8.6 | 7.85 | 8.5 | 8.42 | 7.825 | 7.625 | 8.34 | 8.06 | 8.036 | 8.25 | 7.8 | 7.95 | 4.2715 | 8 | 5.4
| Std Dev. |  | 0.17607 | 0.25156 | 0.19664 | 0.34448 | 0.18619 | 0.16733 | 0.25145 | 0.20686 | 0.32542 | 0.37877 | 0.24887 | 0.15047 | 0.30768 | 0.39707 | 0.19408 | 0.14977 | 0 | 0.18619
| Rel.Std.Dev. |  | 2.363% | 3.306% | 2.906% | 3.979% | 2.367% | 1.969% | 2.992% | 2.667% | 4.321% | 4.572% | 3.088% | 1.872% | 3.768% | 5.08% | 2.421% | 3.488% | 0% | 3.427%
| PDM3 |  | -6.031% | -4.009% | -14.65% | 9.189% | -0.7755% | 7.213% | 5.994% | -2.184% | -5.001% | 4.501% | 1.642% | 1.393% | 3.009% | -1.406% | 1.117% | -45.84% | 0.9063% | -31.47%
| Z-Score (Absolute) |  | 0.84 | 0.52 | 2.21 | 1.59 | 0 | 1.27 | 1.08 | 0.22 | 0.67 | 0.84 | 0.39 | 0.35 | 0.6 | 0.1 | 0.3 | 7.19 | NA | 4.9
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho2O3 in OREAS 471 (Certified Value 1.03 ppm) | [28 empty] | Certified
Ho2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 1.031 | 1.0221 | < 2.3 | 1.0653 | 1.031 | 1.0997 | 1.031 | 1.031 | 1.0539 | 1.031 | 1.0768 | 0.99658 | 1.008 | 1.0424 | 1.031 | 1.0025 | 1.031 | 1.1799 | 1.0424 | 1.0733 | 1.0454 | 1.031 | 0.98427 | 1.157 | 0.97426 | 0.99658 |  | 1
| 1 | 2 | 1.031 | 1.0036 | < 2.3 | 1.008 | 1.0882 | 1.0539 | 1.031 | 1.031 | 1.0424 | 1.031 | 1.134 | 1.0424 | 0.97368 | 1.0195 | 0.95076 | 1.0376 | 1.031 | 0.97368 | 1.031 | 1.0825 | 0.96349 | 0.99658 | 0.98619 | 1.0882 | 0.97373 | 1.0195 |  | #N/A
| 1 | 3 | 1.031 | 1.1428 | < 2.3 | 0.99658 | 1.008 | 1.0653 | 1.031 | 1.031 | 0.98513 | 1.1455 | 1.0882 | 0.96222 | 0.98513 | 1.0424 | 0.97368 | 1.0405 | 1.031 | 0.99658 | 1.0653 | 1.11 | 1.2394 | 1.0653 | 0.96806 | 1.0768 | 0.98408 | 0.99658 |  | 16
| 1 | 4 | 1.031 | 1.0218 | < 2.3 | 1.0539 | 1.0539 | 1.0882 | 1.031 | 1.031 | 0.96222 | 1.031 | 1.0424 | 0.99658 | 0.93931 | 1.0195 | 1.0539 | 0.99418 | 1.031 | 1.134 | 1.031 | 1.063 | 1.1233 | 1.0195 | 0.98781 | 1.1684 | 0.97992 | 0.97368 |  | 1.0281
| 1 | 5 | 1.031 | 1.0371 | < 2.3 | 1.0539 | 1.031 | 1.0653 | 1.031 | 1.031 | 1.134 | 1.031 | 1.1226 | 0.98513 | 0.98513 | 0.97368 | 1.0997 | 1.0563 | 1.031 | 0.99658 | 1.0882 | 1.0848 | 1.2549 | 1.031 | 0.92268 | 1.0882 | 0.98978 | 0.98513 |  | 24
| 1 | 6 | 1.031 | 1.0276 | < 2.3 | 1.0653 | 1.0424 | 1.0882 | 1.031 | 1.031 | 0.96222 | 1.031 | 1.0424 | 0.92785 | 0.97368 | 0.99658 | 1.1226 | 1.0438 | 1.1455 | 1.0424 | 1.031 | 1.0848 | 1.0588 | 1.0653 | 0.98337 | 1.1455 | 0.98559 | 1.0195
| Mean |  | 1.031 | 1.0425 |  | 1.0405 | 1.0424 | 1.0768 | 1.031 | 1.031 | 1.0233 | 1.05 | 1.0844 | 0.98513 | 0.97749 | 1.0157 | 1.0386 | 1.0291 | 1.05 | 1.0539 | 1.0481 | 1.0831 | 1.1142 | 1.0348 | 0.97206 | 1.1207 | 0.98123 | 0.99849
| Median |  | 1.031 | 1.0249 |  | 1.0539 | 1.0367 | 1.0768 | 1.031 | 1.031 | 1.0138 | 1.031 | 1.0825 | 0.99086 | 0.9794 | 1.0195 | 1.0424 | 1.039 | 1.031 | 1.0195 | 1.0367 | 1.0836 | 1.091 | 1.031 | 0.98382 | 1.1169 | 0.982 | 0.99658
| Std Dev. |  | 0 | 0.050326 |  | 0.030235 | 0.027107 | 0.017746 | 0 | 0 | 0.067055 | 0.046765 | 0.038789 | 0.038336 | 0.022525 | 0.026783 | 0.067833 | 0.024817 | 0.046765 | 0.084177 | 0.023754 | 0.015669 | 0.11497 | 0.026783 | 0.025215 | 0.040607 | 0.0064288 | 0.018352
| Rel.Std.Dev. |  | 0% | 4.827% |  | 2.906% | 2.6% | 1.648% | 0% | 0% | 6.553% | 4.454% | 3.577% | 3.891% | 2.304% | 2.637% | 6.531% | 2.411% | 4.454% | 7.987% | 2.266% | 1.447% | 10.32% | 2.588% | 2.594% | 3.623% | 0.6552% | 1.838%
| PDM3 |  | 0.2747% | 1.398% |  | 1.203% | 1.389% | 4.731% | 0.2747% | 0.2747% | -0.4681% | 2.132% | 5.474% | -4.182% | -4.925% | -1.211% | 1.017% | 0.09825% | 2.132% | 2.503% | 1.946% | 5.344% | 8.375% | 0.646% | -5.453% | 9.002% | -4.562% | -2.882%
| Z-Score (Absolute) |  | 0.38 | 0.13 | 4.68 | 0.04 | 0.13 | 1.64 | 0.38 | 0.38 | 0.72 | 0.46 | 1.98 | 2.4 | 2.74 | 1.05 | 0.04 | 0.46 | 0.46 | 0.63 | 0.38 | 1.92 | 3.3 | 0.21 | 2.98 | 3.58 | 2.57 | 1.81
| Analytical results for In in OREAS 471 (Certified Value < 0.1 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 04 | 06 | 08 | 09 | 10 | 17 | 18 | 21 | 27 | [17 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES/MS | [17 empty] | 2
2
| 1 | 1 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | < 0.1 | < 0.2 | 0.0569 | < 0.1 | < 0.1 | 0.089 | [17 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | < 0.1 | < 0.2 | 0.0615 | < 0.1 | < 0.1 | 0.09 | [17 empty] | 7
| 1 | 3 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | < 0.1 | < 0.2 | 0.0569 | < 0.1 | < 0.1 | 0.092 | [17 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | < 0.1 | < 0.2 | 0.0488 | < 0.1 | < 0.1 | 0.087 | [17 empty] | < 0.1
| 1 | 5 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | < 0.1 | < 0.2 | 0.0533 | < 0.1 | < 0.1 | 0.089 | [17 empty] | 25
| 1 | 6 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | < 0.1 | < 0.2 | 0.0527 | < 0.1 | < 0.1 | 0.094
| Mean | [7 empty] | 0.055017 |  |  | 0.090167
| Median | [7 empty] | 0.0551 |  |  | 0.0895
| Std Dev. | [7 empty] | 0.0043838 |  |  | 0.0024833
| Rel.Std.Dev. | [7 empty] | 7.968% |  |  | 2.754%
PDM3
| Z-Score (Absolute) |  | 0.82 | 0.82 | 0.82 | 0.67 | 0.67 | 0.82 | 0.52 | 0.67 | 0.67 | 0.52
| Analytical results for K2O in OREAS 471 (Certified Value 3.28 wt.%) | [28 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 23 | 27 | [12 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 3.2 | 3.28 | 2.8679 | 3.313 | 3.224 | 3.252 | 3.264 | 3.252 | 3.49 | 3.24 | 3.3392 | 3.204 | 3.34 | 3.4813 | 3.24 | [12 empty] | 1
| 1 | 2 | 3.22 | 3.2542 | 3.0743 | 3.096 | 3.2627 | 3.192 | 3.216 | 3.228 | 3.37 | 3.3 | 3.4683 | 3.204 | 3.33 | 3.3163 | 3.264 | [12 empty] | #N/A
| 1 | 3 | 3.18 | 3.211 | 2.9584 | 3.289 | 3.234 | 3.216 | 3.289 | 3.252 | 3.25 | 3.27 | 3.353 | 3.216 | 3.35 | 3.4656 | 3.228 | [12 empty] | 16
| 1 | 4 | 3.2 | 3.2384 | 2.7897 | 3.216 | 3.2423 | 3.096 | 3.252 | 3.264 | 3.25 | 3.26 | 3.3163 | 3.156 | 3.34 | 3.4548 | 3.24 | [12 empty] | 3.2756
| 1 | 5 | 3.21 | 3.2502 | 2.8837 | 3.313 | 3.2896 | 3.24 | 3.277 | 3.228 | 3.37 | 3.32 | 3.4636 | 3.216 | 3.29 | 3.4717 | 3.264 | [12 empty] | 26
| 1 | 6 | 3.21 | 3.2222 | 2.8953 | 3.168 | 3.3118 | 3.12 | 3.228 | 3.301 | 3.37 | 3.26 | 3.3951 | 3.156 | 3.32 | 3.7331 | 3.252
| Mean |  | 3.2033 | 3.2427 | 2.9115 | 3.2325 | 3.2607 | 3.186 | 3.2543 | 3.2542 | 3.35 | 3.275 | 3.3893 | 3.192 | 3.3283 | 3.4871 | 3.248
| Median |  | 3.205 | 3.2443 | 2.8895 | 3.2525 | 3.2525 | 3.204 | 3.258 | 3.252 | 3.37 | 3.265 | 3.374 | 3.204 | 3.335 | 3.4686 | 3.246
| Std Dev. |  | 0.013663 | 0.024618 | 0.096388 | 0.088564 | 0.034215 | 0.064287 | 0.028204 | 0.027088 | 0.090333 | 0.029496 | 0.064733 | 0.028397 | 0.02137 | 0.13525 | 0.014533
| Rel.Std.Dev. |  | 0.4265% | 0.7592% | 3.311% | 2.74% | 1.049% | 2.018% | 0.8667% | 0.8324% | 2.696% | 0.9006% | 1.91% | 0.8896% | 0.6421% | 3.879% | 0.4474%
| PDM3 |  | -2.205% | -1.005% | -11.11% | -1.315% | -0.4531% | -2.735% | -0.6484% | -0.6535% | 2.272% | -0.01746% | 3.47% | -2.551% | 1.611% | 6.459% | -0.8417%
| Z-Score (Absolute) |  | 0.67 | 0.16 | 4.53 | 0.29 | 0.08 | 0.9 | 0 | 0 | 1.3 | 0.27 | 1.78 | 0.82 | 0.98 | 3.07 | 0.08
| Analytical results for La2O3 in OREAS 471 (Certified Value 124 ppm) | [28 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 0
0
| 1 | 1 | 125.4896 | 129.3086 | 119.6256 | 126.6624 | 117.6318 | 125.4896 | 121.9712 | 122.0885 | 124.3168 | 125.8414 | 125.4896 | 119.0392 | 126.6624 | 127.2488 | 116.8109 | 119.7051 | 127.8352 | 124.4341 | 124.4341 | 131.0135 | 125.0339 | 120.7984 | 121.268 | 129.39 | 126.6416 | 119.6256 | 145.66 | 1
| 1 | 2 | 123.144 | 124.4597 | 126.6624 | 121.9712 | 120.9157 | 121.9712 | 123.144 | 125.1378 | 133.6992 | 125.8414 | 123.7304 | 120.7984 | 126.076 | 123.144 | 119.0392 | 123.772 | 126.6624 | 126.5451 | 123.965 | 131.4263 | 114.9293 | 118.4528 | 124.422 | 129.48 | 125.1926 | 120.6811 | 139.25 | #N/A
| 1 | 3 | 126.6624 | 123.0361 | 117.28 | 124.3168 | 120.4466 | 125.4896 | 124.3168 | 123.0267 | 120.7984 | 126.4278 | 120.7984 | 120.212 | 117.8664 | 119.6256 | 119.0392 | 118.1249 | 126.6624 | 124.7859 | 127.9525 | 132.9885 | 115.5322 | 121.9712 | 126.217 | 129.2 | 125.4533 | 119.7429 | 138.54 | 16
| 1 | 4 | 126.6624 | 137.8523 | 111.416 | 124.3168 | 118.4528 | 120.7984 | 124.3168 | 125.0205 | 117.1627 | 128.0698 | 123.144 | 119.0392 | 125.4896 | 123.144 | 121.9712 | 118.5752 | 124.3168 | 123.6131 | 123.3786 | 131.3665 | 134.9773 | 118.4528 | 123.789 | 129.14 | 124.4331 | 119.2738 | 59 | 123.674
| 1 | 5 | 126.6624 | 127.5041 | 118.4528 | 125.4896 | 120.3293 | 124.3168 | 126.6624 | 126.3106 | 140.736 | 127.3661 | 124.9032 | 120.212 | 122.5576 | 117.8664 | 119.0392 | 124.6595 | 126.6624 | 123.6131 | 130.1808 | 130.7555 | 121.4382 | 119.6256 | 127.284 | 129.31 | 125.1359 | 119.5083 | 139.4 | 27
| 1 | 6 | 123.144 | 128.8976 | 119.6256 | 127.8352 | 119.7429 | 119.6256 | 123.144 | 123.4958 | 124.3168 | 125.8414 | 119.0392 | 120.212 | 122.5576 | 116.5763 | 121.3848 | 120.4674 | 125.4896 | 123.2613 | 125.8414 | 131.3747 | 123.9115 | 121.9712 | 124.95 | 130.26 | 126.0009 | 120.212 | 139.65
| Mean |  | 125.2941 | 128.5097 | 118.8437 | 125.0987 | 119.5865 | 122.9485 | 123.9259 | 124.18 | 126.8383 | 126.5647 | 122.8508 | 119.9188 | 123.5349 | 121.2675 | 119.5474 | 120.884 | 126.2715 | 124.3754 | 125.9587 | 131.4875 | 122.6371 | 120.212 | 124.655 | 129.4633 | 125.4762 | 119.8406 | 126.9167
| Median |  | 126.076 | 128.2009 | 119.0392 | 124.9032 | 120.0361 | 123.144 | 123.7304 | 124.2582 | 124.3168 | 126.1346 | 123.4372 | 120.212 | 124.0236 | 121.3848 | 119.0392 | 120.0863 | 126.6624 | 124.0236 | 125.1378 | 131.3706 | 122.6749 | 120.212 | 124.686 | 129.35 | 125.3229 | 119.6842 | 139.325
| Std Dev. |  | 1.7263 | 5.2048 | 4.9015 | 2.0538 | 1.2797 | 2.5062 | 1.6023 | 1.5707 | 8.7525 | 0.9482 | 2.481 | 0.71819 | 3.2894 | 3.9772 | 1.8716 | 2.724 | 1.2113 | 1.2069 | 2.6371 | 0.78033 | 7.3618 | 1.6166 | 2.084 | 0.40932 | 0.76399 | 0.51612 | 33.3742
| Rel.Std.Dev. |  | 1.378% | 4.05% | 4.124% | 1.642% | 1.07% | 2.038% | 1.293% | 1.265% | 6.901% | 0.7492% | 2.019% | 0.5989% | 2.663% | 3.28% | 1.566% | 2.253% | 0.9593% | 0.9704% | 2.094% | 0.5935% | 6.003% | 1.345% | 1.672% | 0.3162% | 0.6089% | 0.4307% | 26.3%
| PDM3 |  | 1.31% | 3.91% | -3.906% | 1.152% | -3.305% | -0.5866% | 0.2037% | 0.4092% | 2.559% | 2.337% | -0.6656% | -3.036% | -0.1124% | -1.946% | -3.337% | -2.256% | 2.1% | 0.5672% | 1.847% | 6.318% | -0.8384% | -2.799% | 0.7932% | 4.681% | 1.457% | -3.1% | 2.622%
| Z-Score (Absolute) |  | 0.32 | 1.22 | 1.51 | 0.26 | 1.3 | 0.35 | 0.07 | 0 | 0.75 | 0.67 | 0.38 | 1.2 | 0.18 | 0.82 | 1.31 | 0.93 | 0.59 | 0.06 | 0.5 | 2.07 | 0.44 | 1.12 | 0.13 | 1.49 | 0.37 | 1.23 | 0.77
| Analytical results for Li in OREAS 471 (Certified Value 16.3 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 03 | 04 | 06 | 08 | 09 | 10 | 18 | 26 | 27 | [18 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [18 empty] | 1
2
| 1 | 1 | 24 | 15 | 17 | 17 | 16 | 20 | < 100 | 14.65 | 16.9 | [18 empty] | 1
| 1 | 2 | 23 | 15 | 18 | 16 | 16 | 15 | < 100 | 13.52 | 17 | [18 empty] | #N/A
| 1 | 3 | 21 | 15 | 19 | 17 | 15 | 15 | < 100 | 15.39 | 17 | [18 empty] | 16
| 1 | 4 | 20 | 15 | 16 | 17 | 17 | 15 | < 100 | 14.95 | 16.8 | [18 empty] | 16.2893
| 1 | 5 | 20 | 15 | 18 | 17 | 16 | 15 | < 100 | 14.75 | 16.9 | [18 empty] | 28
| 1 | 6 | 21 | 15 | 19 | 17 | 16 | 20 | < 100 | 15.39 | 16.9
| Mean |  | 21.5 | 15 | 17.8333 | 16.8333 | 16 | 16.6667 |  | 14.775 | 16.9167
| Median |  | 21 | 15 | 18 | 17 | 16 | 15 |  | 14.85 | 16.9
| Std Dev. |  | 1.6432 | 0 | 1.169 | 0.40825 | 0.63246 | 2.582 |  | 0.68969 | 0.075277
| Rel.Std.Dev. |  | 7.643% | 0% | 6.555% | 2.425% | 3.953% | 15.49% |  | 4.668% | 0.445%
| PDM3 |  | 31.99% | -7.915% | 9.479% | 3.34% | -1.776% | 2.317% |  | -9.296% | 3.851%
| Z-Score (Absolute) |  | 3.15 | 1.24 | 0.67 | 0 | 0.56 | 0.11 | 22.36 | 1.39 | 0.06
| Analytical results for Lu2O3 in OREAS 471 (Certified Value 0.37 ppm) | [28 empty] | Certified
Lu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.38661 | 0.40277 | < 2.3 | 0.37524 | 0.39798 | 0.40936 | 0.36387 | 0.34113 | 0.40936 | 0.34113 | 0.3525 | 0.37524 | 0.38661 | 0.4321 | 0.34113 | 0.37274 | 0.45484 | 0.4321 | 0.3525 | 0.36615 | 0.36076 | 0.36387 | 0.32434 | 0.42073 | 0.3468 | 0.31839 |  | 1
| 1 | 2 | 0.38661 | 0.3969 | < 2.3 | 0.34113 | 0.42073 | 0.42073 | 0.3525 | 0.45484 | 0.4321 | 0.34113 | 0.40936 | 0.32976 | 0.32976 | 0.39798 | 0.36387 | 0.37433 | 0.34113 | 0.39798 | 0.3525 | 0.37069 | 0.32595 | 0.3525 | 0.33752 | 0.38661 | 0.35585 | 0.32976 |  | #N/A
| 1 | 3 | 0.37524 | 0.43637 | < 2.3 | 0.36387 | 0.3525 | 0.46621 | 0.38661 | 0.34113 | 0.39798 | 0.34113 | 0.39798 | 0.32976 | 0.42073 | 0.42073 | 0.32976 | 0.37513 | 0.34113 | 0.34113 | 0.36387 | 0.38207 | 0.39668 | 0.37524 | 0.3327 | 0.38661 | 0.35237 | 0.32976 |  | 16
| 1 | 4 | 0.37524 | 0.40421 | < 2.3 | 0.37524 | 0.31839 | 0.42073 | 0.36387 | 0.34113 | 0.3525 | 0.34113 | 0.4321 | 0.3525 | 0.3525 | 0.38661 | 0.3525 | 0.37627 | 0.34113 | 0.4321 | 0.3525 | 0.37752 | 0.37192 | 0.36387 | 0.35785 | 0.36387 | 0.35529 | 0.31839 |  | 0.36775
| 1 | 5 | 0.38661 | 0.4316 | < 2.3 | 0.36387 | 0.3525 | 0.47758 | 0.37524 | 0.45484 | 0.40936 | 0.45484 | 0.36387 | 0.36387 | 0.36387 | 0.31839 | 0.32976 | 0.38752 | 0.34113 | 0.29565 | 0.37524 | 0.37069 | 0.4394 | 0.36387 | 0.35468 | 0.3525 | 0.344 | 0.32976 |  | 29
| 1 | 6 | 0.38661 | 0.43451 | < 2.3 | 0.38661 | 0.36387 | 0.46621 | 0.3525 | 0.34113 | 0.40936 | 0.45484 | 0.34113 | 0.40936 | 0.34113 | 0.34113 | 0.30702 | 0.37286 | 0.34113 | 0.4321 | 0.37524 | 0.36956 | 0.36051 | 0.37524 | 0.33595 | 0.36387 | 0.345 | 0.32976
| Mean |  | 0.38282 | 0.41773 |  | 0.36766 | 0.36766 | 0.44347 | 0.36577 | 0.37903 | 0.40178 | 0.37903 | 0.38282 | 0.36008 | 0.36577 | 0.38282 | 0.33734 | 0.37647 | 0.36008 | 0.38851 | 0.36198 | 0.37278 | 0.37587 | 0.36577 | 0.34051 | 0.37903 | 0.34989 | 0.32597
| Median |  | 0.38661 | 0.4179 |  | 0.36956 | 0.35819 | 0.44347 | 0.36387 | 0.34113 | 0.40936 | 0.34113 | 0.38093 | 0.35819 | 0.35819 | 0.3923 | 0.33544 | 0.37473 | 0.34113 | 0.41504 | 0.35819 | 0.37069 | 0.36634 | 0.36387 | 0.33673 | 0.37524 | 0.34959 | 0.32976
| Std Dev. |  | 0.005872 | 0.018231 |  | 0.015536 | 0.036435 | 0.029652 | 0.013293 | 0.05872 | 0.026586 | 0.05872 | 0.035718 | 0.030228 | 0.033282 | 0.04472 | 0.019913 | 0.0055778 | 0.046422 | 0.05772 | 0.01118 | 0.0058609 | 0.038579 | 0.0085598 | 0.013068 | 0.024564 | 0.0052722 | 0.005872
| Rel.Std.Dev. |  | 1.534% | 4.364% |  | 4.226% | 9.91% | 6.686% | 3.634% | 15.49% | 6.617% | 15.49% | 9.33% | 8.395% | 9.099% | 11.68% | 5.903% | 1.482% | 12.89% | 14.86% | 3.089% | 1.572% | 10.26% | 2.34% | 3.838% | 6.481% | 1.507% | 1.801%
| PDM3 |  | 4.099% | 13.59% |  | -0.02383% | -0.02383% | 20.59% | -0.5392% | 3.068% | 9.252% | 3.068% | 4.099% | -2.085% | -0.5392% | 4.099% | -8.269% | 2.373% | -2.085% | 5.645% | -1.57% | 1.368% | 2.208% | -0.5392% | -7.408% | 3.068% | -4.857% | -11.36%
| Z-Score (Absolute) |  | 0.67 | 3.02 | NA | 0.34 | 0.34 | 4.75 | 0.47 | 0.42 | 1.95 | 0.42 | 0.67 | 0.85 | 0.47 | 0.67 | 2.38 | 0.25 | 0.85 | 1.06 | 0.73 | 0 | 0.21 | 0.47 | 2.17 | 0.42 | 1.54 | 3.14
Results from laboratory 1.03 were removed due to their systematic bias.
| Analytical results for MgO in OREAS 471 (Certified Value 1.3 wt.%) | [28 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 23 | 24 | 26 | 27 | [10 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 1.26 | 1.3293 | 1.2989 | 1.31 | 1.2821 | 1.526 | 1.26 | 1.31 | 1.327 | 1.29 | 1.3174 | 1.327 | 1.33 | 1.2801 | 1.2926 | 1.4053 | 1.31 | [10 empty] | 1
| 1 | 2 | 1.26 | 1.2973 | 1.3541 | 1.293 | 1.2969 | 1.592 | 1.26 | 1.31 | 1.327 | 1.31 | 1.3395 | 1.327 | 1.34 | 1.2503 | 1.2894 | 1.3927 | 1.26 | [10 empty] | #N/A
| 1 | 3 | 1.25 | 1.2901 | 1.2664 | 1.31 | 1.2718 | 1.526 | 1.26 | 1.327 | 1.293 | 1.29 | 1.3188 | 1.327 | 1.33 | 1.2818 | 1.287 | 1.4133 | 1.35 | [10 empty] | 16
| 1 | 4 | 1.25 | 1.3046 | 1.2087 | 1.31 | 1.3101 | 1.443 | 1.26 | 1.327 | 1.21 | 1.29 | 1.2959 | 1.31 | 1.32 | 1.237 | 1.2884 | 1.3987 | 1.31 | [10 empty] | 1.2998
| 1 | 5 | 1.26 | 1.2911 | 1.2354 | 1.327 | 1.2949 | 1.492 | 1.26 | 1.31 | 1.327 | 1.31 | 1.348 | 1.327 | 1.34 | 1.2851 | 1.2875 | 1.3892 | 1.35 | [10 empty] | 30
| 1 | 6 | 1.25 | 1.2979 | 1.3009 | 1.31 | 1.3133 | 1.575 | 1.26 | 1.343 | 1.293 | 1.29 | 1.3379 | 1.31 | 1.34 | 1.2834 | 1.2865 | 1.4067 | 1.27
| Mean |  | 1.255 | 1.3017 | 1.2774 | 1.31 | 1.2949 | 1.5257 | 1.26 | 1.3212 | 1.2962 | 1.2967 | 1.3262 | 1.3213 | 1.3333 | 1.2696 | 1.2885 | 1.401 | 1.3083
| Median |  | 1.255 | 1.2976 | 1.2826 | 1.31 | 1.2959 | 1.526 | 1.26 | 1.3185 | 1.31 | 1.29 | 1.3283 | 1.327 | 1.335 | 1.281 | 1.2879 | 1.402 | 1.31
| Std Dev. |  | 0.0054772 | 0.014486 | 0.051975 | 0.010752 | 0.015934 | 0.054408 | 0 | 0.013556 | 0.04538 | 0.010328 | 0.01918 | 0.0087788 | 0.008165 | 0.020616 | 0.0022213 | 0.0091263 | 0.038166
| Rel.Std.Dev. |  | 0.4364% | 1.113% | 4.069% | 0.8207% | 1.231% | 3.566% | 0% | 1.026% | 3.501% | 0.7965% | 1.446% | 0.6644% | 0.6124% | 1.624% | 0.1724% | 0.6514% | 2.917%
| PDM3 |  | -3.449% | 0.1439% | -1.728% | 0.7821% | -0.3827% | 17.37% | -3.065% | 1.641% | -0.2821% | -0.2437% | 2.032% | 1.654% | 2.577% | -2.325% | -0.8693% | 7.781% | 0.6539%
| Z-Score (Absolute) |  | 1.63 | 0 | 0.85 | 0.29 | 0.24 | 7.81 | 1.45 | 0.67 | 0.19 | 0.18 | 0.86 | 0.67 | 1.1 | 1.12 | 0.46 | 3.46 | 0.23
| Analytical results for MnO in OREAS 471 (Certified Value 0.05 wt.%) | [28 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 21 | 23 | 24 | 27 | [10 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | BF*OES | PF*OES | PF*OES/MS | [10 empty] | 3
3
| 1 | 1 | 0.051 | 0.053045 | 0.05165 | 0.05113 | 0.05 | 0.05552 | 0.0491 | < 0.26 | 0.05281 | 0.05 | 0.049 | < 0.26 | 0.05 | 0.047517 | 0.0516 | 0.048337 | 0.05023 | [10 empty] | 1
| 1 | 2 | 0.052 | 0.053876 | 0.05359 | 0.04868 | 0.05 | 0.05359 | 0.0491 | < 0.26 | 0.05268 | 0.05 | 0.0503 | < 0.26 | 0.05 | 0.049002 | 0.0504 | 0.048601 | 0.05049 | [10 empty] | #N/A
| 1 | 3 | 0.051 | 0.051615 | 0.05281 | 0.05036 | 0.05 | 0.05707 | 0.0491 | < 0.26 | 0.05281 | 0.05 | 0.0496 | < 0.26 | 0.05 | 0.049002 | 0.0529 | 0.048856 | 0.0501 | [10 empty] | 16
| 1 | 4 | 0.051 | 0.054306 | 0.04997 | 0.0501 | 0.05 | 0.05359 | 0.0491 | < 0.26 | 0.04816 | 0.05 | 0.0494 | < 0.26 | 0.05 | 0.050487 | 0.0504 | 0.048469 | 0.04932 | [10 empty] | 0.050498
| 1 | 5 | 0.051 | 0.052069 | 0.05062 | 0.05126 | 0.05 | 0.0563 | 0.0491 | < 0.26 | 0.05268 | 0.05 | 0.0514 | < 0.26 | 0.05 | 0.049002 | 0.0516 | 0.04811 | 0.04997 | [10 empty] | 31
| 1 | 6 | 0.051 | 0.053542 | 0.05307 | 0.04971 | 0.05 | 0.05423 | 0.0491 | < 0.26 | 0.05217 | 0.05 | 0.0505 | < 0.26 | 0.05 | 0.044547 | 0.0504 | 0.048625 | 0.05062
| Mean |  | 0.051167 | 0.053076 | 0.051952 | 0.050207 | 0.05 | 0.05505 | 0.0491 |  | 0.051885 | 0.05 | 0.050033 |  | 0.05 | 0.048259 | 0.051217 | 0.0485 | 0.050122
| Median |  | 0.051 | 0.053294 | 0.05223 | 0.05023 | 0.05 | 0.054875 | 0.0491 |  | 0.05268 | 0.05 | 0.04995 |  | 0.05 | 0.049002 | 0.051 | 0.048535 | 0.050165
| Std Dev. |  | 0.00040825 | 0.0010504 | 0.0014466 | 0.00095649 | 7.6012e-18 | 0.0014697 | 0 |  | 0.0018402 | 7.6012e-18 | 0.00087331 |  | 7.6012e-18 | 0.0020469 | 0.0010128 | 0.0002577 | 0.00046084
| Rel.Std.Dev. |  | 0.7979% | 1.979% | 2.784% | 1.905% | 1.52e-14% | 2.67% | 0% |  | 3.547% | 1.52e-14% | 1.745% |  | 1.52e-14% | 4.242% | 1.977% | 0.5313% | 0.9194%
| PDM3 |  | 1.324% | 5.105% | 2.879% | -0.5769% | -0.9862% | 9.014% | -2.768% |  | 2.747% | -0.9862% | -0.9202% |  | -0.9862% | -4.433% | 1.423% | -3.957% | -0.7452%
| Z-Score (Absolute) |  | 0.57 | 1.7 | 1.03 | 0 | 0.12 | 2.86 | 0.67 | 46.65 | 0.99 | 0.12 | 0.1 | 46.65 | 0.12 | 1.15 | 0.6 | 1.01 | 0.05
| Analytical results for Mo in OREAS 471 (Certified Value 2.33 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 10 | 11 | 17 | 18 | 21 | 23 | 27 | [14 empty] | ppm
|  |  | BF*OES/MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*OES/MS | [14 empty] | 2
2
| 1 | 1 | 2 | 4 | 2 | 2 | 3.8 | 3 | 6 | 2 | 1.8323 | < 50 | 2.103 | < 2 | 1.99 | [14 empty] | 1
| 1 | 2 | 2 | 3 | 2 | 3 | 3.67 | 4 | 7 | 2 | 1.7564 | < 50 | 2.159 | < 2 | 1.98 | [14 empty] | #N/A
| 1 | 3 | 2 | 3 | 2 | 3 | 3.72 | 3 | 6 | 2 | 1.898 | < 50 | 2.204 | < 2 | 1.97 | [14 empty] | 16
| 1 | 4 | 2 | 3 | 2 | 2 | 3.9 | 3 | 6 | 2 | 2.1527 | < 50 | 2.15 | 2 | 2.01 | [14 empty] | 2.3297
| 1 | 5 | 2 | 3 | 3 | 3 | 4.03 | 3 | 6 | 2 | 1.9059 | < 50 | 2.089 | 2 | 1.99 | [14 empty] | 32
| 1 | 6 | 2 | 4 | 2 | 2 | 4.13 | 4 | 6 | 2 | 1.8556 | < 50 | 2.037 | 5 | 2
| Mean |  | 2 | 3.3333 | 2.1667 | 2.5 | 3.875 | 3.3333 | 6.1667 | 2 | 1.9002 |  | 2.1237 | 3 | 1.99
| Median |  | 2 | 3 | 2 | 2.5 | 3.85 | 3 | 6 | 2 | 1.8768 |  | 2.1265 | 2 | 1.99
| Std Dev. |  | 0 | 0.5164 | 0.40825 | 0.54772 | 0.17964 | 0.5164 | 0.40825 | 0 | 0.13496 |  | 0.05922 | 1.7321 | 0.014142
| Rel.Std.Dev. |  | 0% | 15.49% | 18.84% | 21.91% | 4.636% | 15.49% | 6.62% | 0% | 7.103% |  | 2.789% | 57.74% | 0.7107%
| PDM3 |  | -14.15% | 43.08% | -6.997% | 7.312% | 66.33% | 43.08% | 164.7% | -14.15% | -18.44% |  | -8.842% | 28.77% | -14.58%
| Z-Score (Absolute) |  | 0.42 | 2.95 | 0 | 0.84 | 4.32 | 2.95 | 10.12 | 0.42 | 0.67 | 57.77 | 0.11 | 0.42 | 0.45
| Analytical results for Na2O in OREAS 471 (Certified Value 1.22 wt.%) | [28 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 01 | 02 | 04 | 06 | 07 | 11 | 17 | 20 | 23 | 27 | [17 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | BF*OES/MS | BF*OES | PF*MS | BF*OES | BF*MS | BF*OES/MS | BF*OES | PF*OES/MS | [17 empty] | 2
3
| 1 | 1 | 1.2 | 1.2959 | 1.24 | 1.18 | 1.254 | 1.21 | 0.00012709 | 1.19 | 0.9301 | 1.34 | [17 empty] | 1
| 1 | 2 | 1.2 | 1.2879 | 1.186 | 1.19 | 1.227 | 1.23 | 0.00012918 | 1.19 | 1.0407 | 1.32 | [17 empty] | #N/A
| 1 | 3 | 1.19 | 1.2588 | 1.227 | 1.2 | 1.254 | 1.21 | 0.00012739 | 1.2 | 1.0541 | 1.38 | [17 empty] | 16
| 1 | 4 | 1.19 | 1.2708 | 1.213 | 1.19 | 1.186 | 1.21 | 0.00012443 | 1.19 | 1.0771 | 1.24 | [17 empty] | 1.2248
| 1 | 5 | 1.2 | 1.2739 | 1.24 | 1.19 | 1.24 | 1.23 | 0.00012924 | 1.17 | 0.9611 | 1.25 | [17 empty] | 33
| 1 | 6 | 1.19 | 1.2664 | 1.186 | 1.19 | 1.227 | 1.21 | 0.00012821 | 1.18 | 1.0973 | 1.29
| Mean |  | 1.195 | 1.2756 | 1.2153 | 1.19 | 1.2313 | 1.2167 | 0.00012759 | 1.1867 | 1.0267 | 1.3033
| Median |  | 1.195 | 1.2723 | 1.22 | 1.19 | 1.2335 | 1.21 | 0.0001278 | 1.19 | 1.0474 | 1.305
| Std Dev. |  | 0.0054772 | 0.013846 | 0.024817 | 0.0063246 | 0.02528 | 0.010328 | 1.7838e-06 | 0.010328 | 0.066492 | 0.053914
| Rel.Std.Dev. |  | 0.4583% | 1.085% | 2.042% | 0.5315% | 2.053% | 0.8489% | 1.398% | 0.8703% | 6.476% | 4.137%
| PDM3 |  | -2.435% | 4.146% | -0.7744% | -2.843% | 0.5319% | -0.6656% | -99.99% | -3.115% | -16.17% | 6.41%
| Z-Score (Absolute) |  | 0.54 | 1.59 | 0 | 0.67 | 0.42 | 0.03 | NA | 0.76 | 5 | 2.33
Results from laboratory 1.17 were removed due to their systematic bias.
| Analytical results for Nb in OREAS 471 (Certified Value 12.7 ppm) | [28 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 27 | [5 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*OES | PF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*OES | PF*MS | PF*OES/MS | [5 empty] | 1
2
| 1 | 1 | 13 | 12.305 | 8 | 12 | 28 | 34 | 16 | 11 | 13.1 | 14.35 | 13.8 | 12.45 | 14 | 12.85 | 12.3477 | < 10 | 14.7 | 12.8 | 11.865 | < 10 | 10.804 | 11.8 | [5 empty] | 1
| 1 | 2 | 13 | 12.6059 | 10 | 11.5 | 49 | 28 | 17 | 11 | 12.9 | 14 | 12.4 | 11.85 | 13.3 | 13.5 | 12.6139 | < 10 | 14.7 | 12.8 | 12.191 | < 10 | 10.7537 | 11.6 | [5 empty] | #N/A
| 1 | 3 | 12 | 12.3429 | 13 | 12.2 | 35 | 31 | < 10 | 12 | 12.9 | 14.1 | 13.6 | 11.9 | 13.55 | 12.85 | 12.2243 | < 10 | 13.4 | 13 | 12.939 | < 10 | 10.7725 | 11.9 | [5 empty] | 16
| 1 | 4 | 12 | 12.5817 | 8 | 11.7 | 35 | 35 | 15 | 10 | 13.2 | 14 | 13.75 | 12.55 | 13.4 | 13.45 | 11.9924 | < 10 | 13.4 | 13 | 12.042 | 10 | 10.605 | 11.4 | [5 empty] | 12.6901
| 1 | 5 | 13 | 12.7277 | 10 | 11.7 | 56 | 37 | 16 | 12 | 13 | 14.05 | 13.75 | 12.25 | 13.05 | 13.1 | 12.6615 | < 10 | 13.7 | 13.2 | 10.479 | < 10 | 11.1422 | 11.6 | [5 empty] | 34
| 1 | 6 | 14 | 12.2478 | 10 | 11.6 | 63 | 34 | 16 | 11 | 12.9 | 13.4 | 13.85 | 11.95 | 12.85 | 13.7 | 12.3878 | < 10 | 14.6 | 12.9 | 11.901 | < 10 | 11.125 | 11.8
| Mean |  | 12.8333 | 12.4685 | 9.8333 | 11.7833 | 44.3333 | 33.1667 | 16 | 11.1667 | 13 | 13.9833 | 13.525 | 12.1583 | 13.3583 | 13.2417 | 12.3713 |  | 14.0833 | 12.95 | 11.9028 | 10 | 10.8671 | 11.6833
| Median |  | 13 | 12.4623 | 10 | 11.7 | 42 | 34 | 16 | 11 | 12.95 | 14.025 | 13.75 | 12.1 | 13.35 | 13.275 | 12.3678 |  | 14.15 | 12.95 | 11.9715 | 10 | 10.7883 | 11.7
| Std Dev. |  | 0.75277 | 0.19498 | 1.8348 | 0.26394 | 13.7647 | 3.1885 | 0.70711 | 0.75277 | 0.12649 | 0.31411 | 0.55745 | 0.30069 | 0.40177 | 0.35975 | 0.24863 |  | 0.64936 | 0.15166 | 0.8006 |  | 0.21757 | 0.18348
| Rel.Std.Dev. |  | 5.866% | 1.564% | 18.66% | 2.24% | 31.05% | 9.614% | 4.419% | 6.741% | 0.973% | 2.246% | 4.122% | 2.473% | 3.008% | 2.717% | 2.01% |  | 4.611% | 1.171% | 6.726% |  | 2.002% | 1.57%
| PDM3 |  | 1.129% | -1.746% | -22.51% | -7.145% | 249.4% | 161.4% | 26.08% | -12% | 2.442% | 10.19% | 6.579% | -4.19% | 5.266% | 4.347% | -2.512% |  | 10.98% | 2.048% | -6.204% | -21.2% | -14.37% | -7.933%
| Z-Score (Absolute) |  | NA | 0.31 | NA | 0.76 | 20.44 | 13.19 | NA | NA | 0.03 | 0.67 | 0.38 | 0.52 | 0.27 | 0.19 | 0.38 | 5.19 | 0.74 | 0 | 0.68 | NA | 1.36 | 0.83
Results from laboratories 1, 3, 9 and 10 were removed due to their 1 ppm reading resolution. Results from laboratory 23 were removed due to their 10 ppm reading resolution.
| Analytical results for Nd2O3 in OREAS 471 (Certified Value 96 ppm) | [28 empty] | Certified
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 0
1
| 1 | 1 | 96.3446 | 94.6361 | 96.8112 | 97.9776 | 100.427 | 102.6432 | 97.511 | 93.5453 | 95.7614 | 92.2622 | 102.6432 | 94.1285 | 98.3275 | 101.4768 | 95.5282 | 91.8425 | 95.9947 | 95.7614 | 91.9123 | 96.7331 | 102.3237 | 95.1782 | 85.293 | 95.1 | 89.6865 | 91.7957 | 97.68 | 1
| 1 | 2 | 94.8283 | 94.6405 | 99.144 | 93.312 | 101.9434 | 103.8096 | 98.5608 | 96.5779 | 102.06 | 94.7117 | 101.7101 | 96.8112 | 95.0616 | 98.0942 | 96.228 | 94.7357 | 97.0445 | 97.0445 | 93.1954 | 97.2988 | 98.7988 | 93.7786 | 85.1838 | 95.85 | 89.1548 | 94.595 | 94.63 | 11
| 1 | 3 | 97.0445 | 101.5617 | 100.3104 | 96.8112 | 100.3104 | 100.3104 | 95.8781 | 93.5453 | 93.1954 | 94.4784 | 100.8936 | 96.5779 | 91.5624 | 97.2778 | 94.8283 | 91.0193 | 97.3944 | 91.4458 | 94.8283 | 97.595 | 97.4343 | 96.3446 | 86.3924 | 95.53 | 88.8396 | 94.3618 | 98.81 | 16
| 1 | 4 | 97.861 | 95.5029 | 81.648 | 96.8112 | 99.0274 | 99.144 | 95.9947 | 95.2949 | 90.0461 | 95.2949 | 102.4099 | 95.5282 | 94.7117 | 99.144 | 98.5608 | 89.9105 | 98.3275 | 89.6962 | 92.3789 | 96.5406 | 108.2975 | 94.0118 | 85.4324 | 95.92 | 87.271 | 93.4286 | 91.55 | 96.3967
| 1 | 5 | 97.0445 | 99.9495 | 90.9792 | 95.6448 | 101.4768 | 101.4768 | 96.9278 | 95.6448 | 105.4426 | 94.3618 | 102.6432 | 97.2778 | 93.312 | 95.6448 | 95.7614 | 94.4392 | 96.228 | 92.7288 | 97.2778 | 95.4197 | 98.9642 | 94.1285 | 85.2443 | 95.84 | 89.2974 | 92.2622 | 99.24 | 35
| 1 | 6 | 94.4784 | 91.4964 | 94.4784 | 97.9776 | 102.4099 | 99.144 | 96.1114 | 95.5282 | 95.8781 | 96.1114 | 98.5608 | 96.6946 | 91.7957 | 94.7117 | 95.2949 | 92.0544 | 96.1114 | 89.6962 | 93.4286 | 96.2233 | 103.853 | 96.5779 | 85.3468 | 95.47 | 88.5582 | 95.7614 | 97.66
| Mean |  | 96.2669 | 96.2978 | 93.8952 | 96.4224 | 100.9325 | 101.088 | 96.8306 | 95.0227 | 97.0639 | 94.5367 | 101.4768 | 96.1697 | 94.1285 | 97.7249 | 96.0336 | 92.3336 | 96.8501 | 92.7288 | 93.8369 | 96.6351 | 101.6119 | 95.0033 | 85.4821 | 95.6183 | 88.8013 | 93.7008 | 96.595
| Median |  | 96.6946 | 95.0717 | 95.6448 | 96.8112 | 100.9519 | 100.8936 | 96.5196 | 95.4115 | 95.8198 | 94.595 | 102.06 | 96.6362 | 94.0118 | 97.686 | 95.6448 | 91.9484 | 96.6362 | 92.0873 | 93.312 | 96.6368 | 100.6439 | 94.6534 | 85.3199 | 95.685 | 88.9972 | 93.8952 | 97.67
| Std Dev. |  | 1.3435 | 3.748 | 6.8709 | 1.7561 | 1.2468 | 1.9047 | 1.0573 | 1.2248 | 5.7005 | 1.2878 | 1.5805 | 1.1541 | 2.5114 | 2.4427 | 1.3231 | 1.9038 | 0.91447 | 3.0939 | 1.9611 | 0.7781 | 4.0645 | 1.2294 | 0.45402 | 0.31378 | 0.84346 | 1.5004 | 2.9505
| Rel.Std.Dev. |  | 1.396% | 3.892% | 7.318% | 1.821% | 1.235% | 1.884% | 1.092% | 1.289% | 5.873% | 1.362% | 1.557% | 1.2% | 2.668% | 2.5% | 1.378% | 2.062% | 0.9442% | 3.337% | 2.09% | 0.8052% | 4% | 1.294% | 0.5311% | 0.3282% | 0.9498% | 1.601% | 3.054%
| PDM3 |  | -0.1346% | -0.1025% | -2.595% | 0.02671% | 4.705% | 4.867% | 0.4502% | -1.425% | 0.6922% | -1.929% | 5.27% | -0.2355% | -2.353% | 1.378% | -0.3766% | -4.215% | 0.4704% | -3.805% | -2.655% | 0.2473% | 5.41% | -1.445% | -11.32% | -0.8074% | -7.879% | -2.797% | 0.2058%
| Z-Score (Absolute) |  | 0.04 | 0.06 | 0.99 | 0.11 | 2.07 | 2.13 | 0.29 | 0.5 | 0.39 | 0.71 | 2.3 | 0 | 0.89 | 0.67 | 0.06 | 1.66 | 0.3 | 1.49 | 1.01 | 0.2 | 2.36 | 0.51 | 4.63 | 0.24 | 3.19 | 1.07 | 0.18
| Analytical results for Ni in OREAS 471 (Certified Value 31.8 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 10 | 17 | 20 | 21 | 23 | 27 | [15 empty] | ppm
|  |  | BF*OES/MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | [15 empty] | 1
1
| 1 | 1 | 30 | 14 | 31 | N.A. | 47 | 40 | 81 | 33.8197 | 32 | 28.156 | 24 | 32.3 | [15 empty] | 1
| 1 | 2 | 30 | 15 | 32 | N.A. | 43 | 30 | 76 | 36.625 | 28 | 29.306 | 23 | 32.8 | [15 empty] | #N/A
| 1 | 3 | 30 | 16 | 30 | N.A. | 48 | 30 | 88 | 34.9945 | 31 | 30.695 | 33 | 32.9 | [15 empty] | 16
| 1 | 4 | 30 | 16 | 32 | N.A. | 43 | 40 | 111 | 34.3699 | 33 | 32.14 | 37 | 32.2 | [15 empty] | 31.8178
| 1 | 5 | 30 | 20 | 35 | N.A. | 47 | 30 | 72 | 34.3457 | 30 | 31.28 | 34 | 32.4 | [15 empty] | 36
| 1 | 6 | 30 | 16 | 35 | N.A. | 48 | 40 | 74 | 36.8836 | 25 | 28.037 | 26 | 33
| Mean |  | 30 | 16.1667 | 32.5 |  | 46 | 35 | 83.6667 | 35.1731 | 29.8333 | 29.9357 | 29.5 | 32.6
| Median |  | 30 | 16 | 32 |  | 47 | 35 | 78.5 | 34.6822 | 30.5 | 30.0005 | 29.5 | 32.6
| Std Dev. |  | 0 | 2.0412 | 2.0736 |  | 2.3664 | 5.4772 | 14.5694 | 1.2827 | 2.9269 | 1.6977 | 5.8907 | 0.34059
| Rel.Std.Dev. |  | 0% | 12.63% | 6.38% |  | 5.144% | 15.65% | 17.41% | 3.647% | 9.811% | 5.671% | 19.97% | 1.045%
| PDM3 |  | -5.713% | -49.19% | 2.144% |  | 44.57% | 10% | 163% | 10.55% | -6.237% | -5.915% | -7.284% | 2.459%
| Z-Score (Absolute) |  | 0.63 | 4.13 | 0 | NA | 3.41 | 0.63 | 12.94 | 0.68 | 0.67 | 0.65 | 0.76 | 0.03
| Analytical results for P2O5 in OREAS 471 (Certified Value 0.322 wt.%) | [28 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 18 | 20 | 23 | 26 | 27 | [11 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS | [11 empty] | 3
3
| 1 | 1 | 0.35 | 0.32755 | 0.32105 | 0.321 | 0.3277 | 0.42303 | 0.321 | 0.321 | 0.275 | 0.32 | 0.3046 | 0.321 | 0.31 | 0.3185 | 0.32956 | 0.33755 | [11 empty] | 1
| 1 | 2 | 0.33 | 0.31718 | 0.3238 | 0.321 | 0.32793 | 0.42532 | 0.321 | 0.321 | 0.298 | 0.33 | 0.3085 | 0.321 | 0.31 | 0.3139 | 0.32996 | 0.34099 | [11 empty] | #N/A
| 1 | 3 | 0.33 | 0.31945 | 0.3112 | 0.321 | 0.31716 | 0.43586 | 0.344 | 0.344 | 0.275 | 0.32 | 0.3036 | 0.344 | 0.31 | 0.3185 | 0.32905 | 0.33801 | [11 empty] | 16
| 1 | 4 | 0.33 | 0.31482 | 0.29699 | 0.321 | 0.32541 | 0.40836 | 0.321 | 0.321 | 0.252 | 0.32 | 0.295 | 0.321 | 0.31 | 0.3139 | 0.3208 | 0.33641 | [11 empty] | 0.32172
| 1 | 5 | 0.31 | 0.31042 | 0.30799 | 0.344 | 0.34053 | 0.41524 | 0.321 | 0.321 | 0.275 | 0.33 | 0.3083 | 0.321 | 0.32 | 0.3185 | 0.32873 | 0.33824 | [11 empty] | 37
| 1 | 6 | 0.32 | 0.31564 | 0.31578 | 0.321 | 0.33137 | 0.40401 | 0.321 | 0.344 | 0.275 | 0.32 | 0.3089 | 0.321 | 0.31 | 0.3277 | 0.33815 | 0.33962
| Mean |  | 0.32833 | 0.31751 | 0.3128 | 0.32483 | 0.32835 | 0.41864 | 0.32483 | 0.32867 | 0.275 | 0.32333 | 0.30482 | 0.32483 | 0.31167 | 0.3185 | 0.32938 | 0.33847
| Median |  | 0.33 | 0.31641 | 0.31349 | 0.321 | 0.32781 | 0.41914 | 0.321 | 0.321 | 0.275 | 0.32 | 0.30645 | 0.321 | 0.31 | 0.3185 | 0.32931 | 0.33813
| Std Dev. |  | 0.013292 | 0.0057579 | 0.0097324 | 0.0093897 | 0.0076411 | 0.011762 | 0.0093897 | 0.011877 | 0.014546 | 0.005164 | 0.0052966 | 0.0093897 | 0.0040825 | 0.005039 | 0.0055026 | 0.0016141
| Rel.Std.Dev. |  | 4.048% | 1.813% | 3.111% | 2.891% | 2.327% | 2.81% | 2.891% | 3.614% | 5.29% | 1.597% | 1.738% | 2.891% | 1.31% | 1.582% | 1.671% | 0.4769%
| PDM3 |  | 2.055% | -1.31% | -2.773% | 0.9667% | 2.06% | 30.12% | 0.9667% | 2.158% | -14.52% | 0.5005% | -5.255% | 0.9667% | -3.126% | -1.002% | 2.379% | 5.205%
| Z-Score (Absolute) |  | 0.46 | 0.89 | 1.47 | 0 | 0.46 | 11.69 | 0 | 0.47 | 6.17 | 0.16 | 2.47 | 0 | 1.61 | 0.76 | 0.59 | 1.72
| Analytical results for Pb in OREAS 471 (Certified Value 33 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 09 | 10 | 17 | 18 | 21 | 27 | [16 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES/MS | [16 empty] | 1
2
| 1 | 1 | 32 | 34 | 37 | 33 | 30 | 32 | 36 | 34.2348 | 31 | 32.944 | 29.3 | [16 empty] | 1
| 1 | 2 | 33 | 33 | 38 | 30 | 30 | 33 | 39 | 34.9968 | 35 | 31.443 | 30.8 | [16 empty] | #N/A
| 1 | 3 | 32 | 34 | 36 | 30 | 30 | 32 | 33 | 33.7376 | 34 | 31.718 | 29.8 | [16 empty] | 16
| 1 | 4 | 32 | 34 | 38 | 30 | 30 | 33 | 37 | 32.9786 | 32 | 31.196 | 29.8 | [16 empty] | 33.0386
| 1 | 5 | 31 | 34 | 35 | 34 | 30 | 33 | 39 | 34.8901 | 36 | 33.456 | 30.6 | [16 empty] | 38
| 1 | 6 | 33 | 34 | 37 | 35 | 30 | 33 | 35 | 35.8494 | 33 | 32.242 | 30.1
| Mean |  | 32.1667 | 33.8333 | 36.8333 | 32 | 30 | 32.6667 | 36.5 | 34.4479 | 33.5 | 32.1665 | 30.0667
| Median |  | 32 | 34 | 37 | 31.5 | 30 | 33 | 36.5 | 34.5624 | 33.5 | 31.98 | 29.95
| Std Dev. |  | 0.75277 | 0.40825 | 1.169 | 2.2804 | 0 | 0.5164 | 2.3452 | 1.0168 | 1.8708 | 0.88764 | 0.55737
| Rel.Std.Dev. |  | 2.34% | 1.207% | 3.174% | 7.126% | 0% | 1.581% | 6.425% | 2.952% | 5.585% | 2.76% | 1.854%
| PDM3 |  | -2.639% | 2.405% | 11.49% | -3.144% | -9.197% | -1.126% | 10.48% | 4.265% | 1.396% | -2.64% | -8.995%
| Z-Score (Absolute) |  | 0.62 | 0.51 | 2.53 | 0.73 | NA | 0.28 | 2.3 | 0.92 | 0.28 | 0.62 | 2.03
Results from laboratory 8 were removed due to their 10 ppm reading resolution.
| Analytical results for Pr6O11 in OREAS 471 (Certified Value 27.7 ppm) | [28 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 1
2
| 1 | 1 | 27.3053 | 29.0017 | 27.7886 | 28.0302 | 26.3146 | 29.4801 | 27.9094 | 27.1845 | 29.9634 | 28.2719 | 28.9968 | 27.4261 | 29.1176 | 29.1176 | 26.4596 | 28.0167 | 28.5135 | 26.7737 | 27.3416 | 27.3935 | 28.0719 | 28.3444 | 27.8442 | 28.06 | 26.6797 | 27.4986 | 29.19 | 1
| 1 | 2 | 27.4261 | 29.6556 | 28.9968 | 26.5804 | 27.0274 | 29.5042 | 27.547 | 28.0302 | 31.679 | 28.5135 | 29.1176 | 27.0637 | 27.547 | 27.547 | 26.9429 | 28.8423 | 28.6343 | 28.1994 | 26.8945 | 27.4213 | 33.049 | 27.8369 | 27.3984 | 27.74 | 27.2429 | 27.0516 | 30.1 | #N/A
| 1 | 3 | 27.7886 | 28.9911 | 28.9968 | 27.4261 | 26.7375 | 29.4922 | 27.9094 | 27.4261 | 28.7793 | 28.6343 | 28.1511 | 27.547 | 26.5804 | 27.547 | 25.9763 | 27.6772 | 29.7217 | 27.1483 | 28.1269 | 27.82 | 30.562 | 28.4894 | 27.3946 | 27.71 | 26.3875 | 27.0033 | 28.91 | 16
| 1 | 4 | 27.547 | 29.2009 | 24.164 | 27.1845 | 26.5442 | 29.9029 | 27.7886 | 27.547 | 27.1362 | 28.9968 | 28.9968 | 27.6678 | 28.6343 | 27.7886 | 27.6678 | 27.4727 | 28.6343 | 27.9336 | 27.0274 | 27.4322 | 30.1047 | 27.5349 | 27.5106 | 27.94 | 25.9176 | 26.7375 | 28.99 | 27.6542
| 1 | 5 | 27.7886 | 29.7278 | 26.5804 | 27.4261 | 26.9429 | 29.4197 | 27.4261 | 28.0302 | 33.5275 | 28.7552 | 29.2384 | 27.9094 | 27.0637 | 27.3053 | 27.4261 | 28.9369 | 28.5135 | 27.0516 | 28.6464 | 27.1531 | 30.018 | 27.1966 | 27.6155 | 27.88 | 26.4439 | 26.4233 | 28.67 | 39
| 1 | 6 | 27.0637 | 28.4449 | 27.7886 | 28.5135 | 26.9549 | 29.4801 | 27.547 | 27.4261 | 30.1567 | 28.3927 | 28.1511 | 28.2719 | 27.7886 | 26.3388 | 27.4261 | 28.157 | 27.9094 | 26.8583 | 27.2812 | 27.3633 | 32.8136 | 28.5377 | 27.8781 | 27.73 | 26.3738 | 27.402 | 29.16
| Mean |  | 27.4865 | 29.1703 | 27.3859 | 27.5268 | 26.7536 | 29.5465 | 27.6879 | 27.6074 | 30.207 | 28.5941 | 28.7753 | 27.6476 | 27.7886 | 27.6074 | 26.9831 | 28.1838 | 28.6545 | 27.3275 | 27.553 | 27.4306 | 30.7699 | 27.99 | 27.6069 | 27.8433 | 26.5076 | 27.0194 | 29.17
| Median |  | 27.4865 | 29.1013 | 27.7886 | 27.4261 | 26.8402 | 29.4861 | 27.6678 | 27.4865 | 30.06 | 28.5739 | 28.9968 | 27.6074 | 27.6678 | 27.547 | 27.1845 | 28.0868 | 28.5739 | 27.0999 | 27.3114 | 27.4074 | 30.3333 | 28.0906 | 27.5631 | 27.81 | 26.4157 | 27.0274 | 29.075
| Std Dev. |  | 0.28335 | 0.47625 | 1.819 | 0.67234 | 0.27864 | 0.17703 | 0.2081 | 0.34808 | 2.2236 | 0.261 | 0.49176 | 0.41444 | 0.95134 | 0.89684 | 0.65585 | 0.59862 | 0.58901 | 0.59374 | 0.68605 | 0.21663 | 1.8816 | 0.55407 | 0.21341 | 0.14067 | 0.43697 | 0.40302 | 0.49303
| Rel.Std.Dev. |  | 1.031% | 1.633% | 6.642% | 2.442% | 1.041% | 0.5992% | 0.7516% | 1.261% | 7.361% | 0.9128% | 1.709% | 1.499% | 3.423% | 3.249% | 2.431% | 2.124% | 2.056% | 2.173% | 2.49% | 0.7897% | 6.115% | 1.98% | 0.773% | 0.5052% | 1.648% | 1.492% | 1.69%
| PDM3 |  | -0.6061% | 5.483% | -0.9702% | -0.4605% | -3.257% | 6.843% | 0.1221% | -0.1692% | 9.231% | 3.399% | 4.054% | -0.02358% | 0.4861% | -0.1692% | -2.427% | 1.915% | 3.617% | -1.181% | -0.3658% | -0.8085% | 11.27% | 1.214% | -0.1709% | 0.6841% | -4.146% | -2.295% | 5.481%
| Z-Score (Absolute) |  | 0.32 | 3 | 0.52 | 0.24 | 1.76 | 3.74 | 0.08 | 0.08 | 5.04 | 1.86 | 2.22 | 0 | 0.28 | 0.08 | 1.31 | 1.06 | 1.98 | 0.63 | 0.19 | 0.43 | 6.15 | 0.67 | 0.08 | 0.39 | 2.25 | 1.24 | 3
| Analytical results for Rb in OREAS 471 (Certified Value 117 ppm) | [28 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 26 | 27 | [5 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 116 | 122.9522 | 117 | 111 | 113 | 117.2 | 114 | 118.5 | 128 | 117 | 113 | 130.5 | 116 | 111.151 | 129.4 | 118.5 | 116.7 | 117.652 | 83.6 | 105.2651 | 144.3438 | 113.4 | [5 empty] | 1
| 1 | 2 | 117 | 122.1111 | 114 | 114 | 114 | 117.2 | 121 | 120.5 | 125.5 | 119 | 110 | 121 | 114 | 116.2809 | 122.9 | 124.4 | 116.3 | 115.278 | 77.1 | 105.2092 | 138.813 | 113 | [5 empty] | #N/A
| 1 | 3 | 116 | 119.554 | 116 | 113 | 111 | 116.6 | 129 | 121.1 | 122.5 | 121.5 | 107.5 | 121 | 119.5 | 111.6165 | 125.6 | 119.3 | 118.8 | 119.307 | 107 | 105.6062 | 145.035 | 112.6 | [5 empty] | 16
| 1 | 4 | 115 | 123.1151 | 118 | 111 | 113 | 118.7 | 107 | 119 | 123 | 118.5 | 113.5 | 122.5 | 122 | 110.5634 | 123.7 | 118.3 | 118.6 | 115.115 | 118 | 106.4053 | 137.8672 | 115.8 | [5 empty] | 117.4999
| 1 | 5 | 116 | 122.2016 | 116 | 114 | 115 | 118.7 | 120 | 122.1 | 123.5 | 119.5 | 107.5 | 116 | 119 | 115.612 | 121.7 | 119.2 | 121.8 | 117.352 | 116 | 105.5848 | 141.1896 | 114 | [5 empty] | 40
| 1 | 6 | 113 | 122.2299 | 116 | 113 | 118 | 119.7 | 119 | 117.2 | 120 | 121.5 | 110 | 119.5 | 116 | 114.7029 | 124.7 | 116 | 116.8 | 114.531 | 92.9 | 105.8997 | 143.3802 | 114.3
| Mean |  | 115.5 | 122.0273 | 116.1667 | 112.6667 | 114 | 118.0167 | 118.3333 | 119.7333 | 123.75 | 119.5 | 110.25 | 121.75 | 117.75 | 113.3211 | 124.6667 | 119.2833 | 118.1667 | 116.5392 | 99.1 | 105.6617 | 141.7715 | 113.85
| Median |  | 116 | 122.2158 | 116 | 113 | 113.5 | 117.95 | 119.5 | 119.75 | 123.25 | 119.25 | 110 | 121 | 117.5 | 113.1597 | 124.2 | 118.85 | 117.7 | 116.315 | 99.95 | 105.5955 | 142.2849 | 113.7
| Std Dev. |  | 1.3784 | 1.2833 | 1.3292 | 1.3663 | 2.3664 | 1.1923 | 7.3666 | 1.8184 | 2.734 | 1.7607 | 2.5836 | 4.8244 | 2.9283 | 2.4955 | 2.6882 | 2.7766 | 2.0637 | 1.8553 | 17.1343 | 0.44285 | 2.9729 | 1.1415
| Rel.Std.Dev. |  | 1.193% | 1.052% | 1.144% | 1.213% | 2.076% | 1.01% | 6.225% | 1.519% | 2.209% | 1.473% | 2.343% | 3.963% | 2.487% | 2.202% | 2.156% | 2.328% | 1.746% | 1.592% | 17.29% | 0.4191% | 2.097% | 1.003%
| PDM3 |  | -1.702% | 3.853% | -1.135% | -4.113% | -2.979% | 0.4398% | 0.7093% | 1.901% | 5.319% | 1.702% | -6.17% | 3.617% | 0.2129% | -3.556% | 6.099% | 1.518% | 0.5675% | -0.8176% | -15.66% | -10.08% | 20.66% | -3.106%
| Z-Score (Absolute) |  | 0.41 | 0.72 | 0.3 | 0.91 | 0.68 | 0.02 | 0.08 | 0.32 | 1.02 | 0.28 | 1.33 | 0.67 | 0.02 | 0.79 | 1.18 | 0.24 | 0.05 | 0.23 | 3.27 | 2.13 | 4.16 | 0.7
| Analytical results for Re in OREAS 471 (Indicative Value < 0.1 ppm) | [28 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 08 | 09 | 10 | 18 | 27 | [22 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | [22 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.006 | [22 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.006 | [22 empty] | 3
| 1 | 3 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.005 | [22 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.005 | [22 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.006 | [22 empty] | 9
| 1 | 6 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.006
| Mean | [5 empty] | 0.0056667
| Median | [5 empty] | 0.006
| Std Dev. | [5 empty] | 0.0005164
| Rel.Std.Dev. | [5 empty] | 9.113%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 471 (Certified Value 0.036 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 03 | 04 | 07 | 08 | 09 | 10 | 18 | 27 | [19 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | [19 empty] | 3
3
| 1 | 1 | 0.0513 | 0.03 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.03 | [19 empty] | 1
| 1 | 2 | 0.0499 | 0.02 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.03 | [19 empty] | 22
| 1 | 3 | 0.0401 | 0.02 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.03 | [19 empty] | 16
| 1 | 4 | 0.0415 | 0.02 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.03 | [19 empty] | 0.036103
| 1 | 5 | 0.0376 | 0.02 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.03 | [19 empty] | 41
| 1 | 6 | 0.0393 | 0.03 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.05 | 0.03
| Mean |  | 0.043283 | 0.023333 | 0.03 |  |  | 0.04 | 0.05 | 0.03
| Median |  | 0.0408 | 0.02 | 0.03 |  |  | 0.04 | 0.05 | 0.03
| Std Dev. |  | 0.0058229 | 0.005164 | 0 |  |  | 0 | 7.6012e-18 | 0
| Rel.Std.Dev. |  | 13.45% | 22.13% | 0% |  |  | 0% | 1.52e-14% | 0%
| PDM3 |  | 19.89% | -35.37% | -16.9% |  |  | 10.79% | 38.49% | -16.9%
| Z-Score (Absolute) |  | 1.54 | 0.77 | 0 | 0.58 | 0.58 | 1.16 | 2.31 | 0
| Analytical results for Sb in OREAS 471 (Certified Value 0.93 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 18 | 21 | 27 | [16 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES/MS | [16 empty] | 2
2
| 1 | 1 | 0.7 | 1 | 1 | < 5 | 1.2 | 0.9 | 1 | 1 | < 0.5 | 1.724 | 0.93 | [16 empty] | 1
| 1 | 2 | 0.8 | < 1 | 0.9 | < 5 | 1.2 | 0.8 | 1 | 0.9 | 1.6 | 1.738 | 0.96 | [16 empty] | #N/A
| 1 | 3 | 0.8 | 1 | 1 | < 5 | 1.1 | 0.8 | 1 | 1.3 | < 0.5 | 1.785 | 0.94 | [16 empty] | 16
| 1 | 4 | 0.7 | 1 | 0.8 | < 5 | 1.1 | 0.7 | < 1 | 1 | < 0.5 | 1.736 | 0.93 | [16 empty] | 0.93333
| 1 | 5 | 0.8 | 1 | 0.9 | < 5 | 1.1 | 0.9 | 1 | 1 | 1 | 1.69 | 0.92 | [16 empty] | 42
| 1 | 6 | 0.7 | < 1 | 0.8 | < 5 | 1.2 | 0.8 | 2 | 1 | < 0.5 | 1.63 | 0.94
| Mean |  | 0.75 | 1 | 0.9 |  | 1.15 | 0.81667 | 1.2 | 1.0333 | 1.3 | 1.7172 | 0.93667
| Median |  | 0.75 | 1 | 0.9 |  | 1.15 | 0.8 | 1 | 1 | 1.3 | 1.73 | 0.935
| Std Dev. |  | 0.054772 | 0 | 0.089443 |  | 0.054772 | 0.075277 | 0.44721 | 0.13663 | 0.42426 | 0.052484 | 0.013663
| Rel.Std.Dev. |  | 7.303% | 0% | 9.938% |  | 4.763% | 9.218% | 37.27% | 13.22% | 32.64% | 3.056% | 1.459%
| PDM3 |  | -19.64% | 7.143% | -3.571% |  | 23.21% | -12.5% | 28.57% | 10.71% | 39.29% | 83.98% | 0.3571%
| Z-Score (Absolute) |  | 0.67 | NA | 0.13 | 5.65 | 0.77 | 0.43 | NA | 0.35 | 1.22 | 2.82 | 0
Results from laboratories 4 and 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 471 (Certified Value 13.3 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 27 |  |  |  |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*OES | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 13 | 14.7139 | 14 | 12.5 | 11 | 14 | < 20 | < 20 | 15 | 13 | 13.4 | 13.5 | 15.2 | 14.9 | 12.8 | 12.9984 | 15 | 13 | 13 | 13.131 | 14.377 | 12 | 12 |  |  |  |  | 1
| 1 | 2 | 13 | 13.6276 | 15 | 12 | 12 | 14 | < 20 | < 20 | 15 | 13 | 14.2 | 12.6 | 14.4 | 12.7 | 13.2 | 13.2856 | 14 | 12.8 | 13 | 13.388 | 14.2065 | 12 | 12.2 |  |  |  |  | #N/A
| 1 | 3 | 12 | 14.1382 | 14 | 12.8 | 12 | 13 | < 20 | < 20 | 15 | 13 | 13.2 | 13.7 | 14.2 | 13.2 | 13 | 13.0377 | 15 | 12.6 | 13 | 14.897 | 14.7693 | 13 | 12.1 |  |  |  |  | 16
| 1 | 4 | 12 | 13.8491 | 13 | 12.2 | 11 | 14 | < 20 | < 20 | 15 | 13 | 14.8 | 12.4 | 14.3 | 15.3 | 12 | 12.694 | 14 | 12.7 | 13 | 14.7 | 15.354 | 12 | 12 |  |  |  |  | 13.3204
| 1 | 5 | 13 | 13.3606 | 14 | 12.1 | 12 | 14 | < 20 | < 20 | 15 | 13 | 13.6 | 14 | 14 | 12.6 | 13.4 | 13.3544 | 15 | 13.1 | 12 | 14.272 | 14.468 | 12 | 12 |  |  |  |  | 43
| 1 | 6 | 12 | 13.762 | 14 | 12 | 12 | 13 | < 20 | < 20 | 15 | 13 | 13.6 | 13.2 | 14.2 | 12 | 14 | 13.2754 | 14 | 13.1 | 13 | 12.121 | 14.6692 | 12 | 12.1
| Mean |  | 12.5 | 13.9086 | 14 | 12.2667 | 11.6667 | 13.6667 |  |  | 15 | 13 | 13.8 | 13.2333 | 14.3833 | 13.45 | 13.0667 | 13.1076 | 14.5 | 12.8833 | 12.8333 | 13.7515 | 14.6407 | 12.1667 | 12.0667
| Median |  | 12.5 | 13.8056 | 14 | 12.15 | 12 | 14 |  |  | 15 | 13 | 13.6 | 13.35 | 14.25 | 12.95 | 13.1 | 13.1565 | 14.5 | 12.9 | 13 | 13.83 | 14.5686 | 12 | 12.05
| Std Dev. |  | 0.54772 | 0.47022 | 0.63246 | 0.32042 | 0.5164 | 0.5164 |  |  | 0 | 0 | 0.5933 | 0.62823 | 0.42151 | 1.3398 | 0.66533 | 0.24845 | 0.54772 | 0.2137 | 0.40825 | 1.0638 | 0.40349 | 0.40825 | 0.08165
| Rel.Std.Dev. |  | 4.382% | 3.381% | 4.518% | 2.612% | 4.426% | 3.779% |  |  | 0% | 0% | 4.299% | 4.747% | 2.931% | 9.961% | 5.092% | 1.895% | 3.777% | 1.659% | 3.181% | 7.736% | 2.756% | 3.355% | 0.6767%
| PDM3 |  | -6.159% | 4.416% | 5.102% | -7.911% | -12.42% | 2.6% |  |  | 12.61% | -2.405% | 3.601% | -0.6536% | 7.98% | 0.973% | -1.905% | -1.598% | 8.856% | -3.281% | -3.657% | 3.236% | 9.912% | -8.661% | -9.412%
| Z-Score (Absolute) |  | 0.51 | 0.67 | 0.75 | 0.71 | 1.21 | 0.47 | 2.62 | 2.62 | 1.59 | 0.09 | 0.58 | 0.11 | 1.07 | 0.29 | 0.03 | 0 | 1.17 | 0.19 | 0.23 | 0.54 | 1.29 | 0.79 | 0.88
| Analytical results for Se in OREAS 471 (Indicative Value < 20 ppm) | [28 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 04 | 08 | 09 | 18 | 21 | 27 | [21 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [21 empty] | 1
1
| 1 | 1 | < 5 | < 20 | < 20 | < 20 | 2.221 | < 1 | [21 empty] | 1
| 1 | 2 | < 5 | < 20 | < 20 | < 20 | 2.331 | < 1 | [21 empty] | 1
| 1 | 3 | < 5 | < 20 | < 20 | < 20 | 2.337 | < 1 | [21 empty] | 16
| 1 | 4 | < 5 | < 20 | < 20 | < 20 | 2.305 | < 1 | [21 empty] | < 20
| 1 | 5 | < 5 | < 20 | < 20 | < 20 | 2.293 | < 1 | [21 empty] | 7
| 1 | 6 | < 5 | < 20 | < 20 | < 20 | 2.099 | < 1
| Mean | [5 empty] | 2.2643
| Median | [5 empty] | 2.299
| Std Dev. | [5 empty] | 0.091009
| Rel.Std.Dev. | [5 empty] | 4.019%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 0.67 | 0.72 | 1.03
| Analytical results for SiO2 in OREAS 471 (Certified Value 63.34 wt.%) | [28 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 02 | 04 | 06 | 08 | 09 | 10 | 11 | 18 | 20 | 23 | 24 | 27 | [14 empty] | wt.%
|  |  | BF*OES/MS | BF*OES/MS | PF*OES/MS | BF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS | [14 empty] | 2
3
| 1 | 1 | 63.01 | 67.6926 | 62.89 | 63.36 | 61.61 | 66.74 | > 64 | 63.78 | 64.18 | 63.64 | 63.7653 | 62.5339 | 62.97 | [14 empty] | 1
| 1 | 2 | 63.25 | 66.7477 | 63.11 | 64.11 | 61.82 | 66.53 | > 64 | 64.69 | 64.39 | 63.3 | 62.4582 | 62.5009 | 64.44 | [14 empty] | #N/A
| 1 | 3 | 62.43 | 65.2994 | 63.11 | 64.29 | 62.25 | 66.96 | > 64 | 63.95 | 64.39 | 63.42 | 63.3503 | 62.5348 | 63.8 | [14 empty] | 16
| 1 | 4 | 62.85 | 68.3706 | 62.89 | 63.9 | 63.11 | 67.81 | > 64 | 64.12 | 63.11 | 63.39 | 62.839 | 62.6815 | 62.76 | [14 empty] | 63.3435
| 1 | 5 | 62.86 | 65.5229 | 63.11 | 63.89 | 63.32 | 66.53 | > 64 | 65.12 | 64.6 | 63.53 | 63.3075 | 62.6321 | 62.65 | [14 empty] | 44
| 1 | 6 | 63.47 | 64.574 | 62.46 | 64.01 | 63.96 | 68.67 | > 64 | 64.14 | 63.32 | 63.7 | 63.438 | 62.5204 | 63.59
| Mean |  | 62.9783 | 66.3679 | 62.9283 | 63.9267 | 62.6783 | 67.2067 |  | 64.3 | 63.9983 | 63.4967 | 63.193 | 62.5673 | 63.3683
| Median |  | 62.935 | 66.1353 | 63 | 63.955 | 62.68 | 66.85 |  | 64.13 | 64.285 | 63.475 | 63.3289 | 62.5344 | 63.28
| Std Dev. |  | 0.35991 | 1.4821 | 0.25349 | 0.31488 | 0.92761 | 0.8602 |  | 0.50505 | 0.62467 | 0.15423 | 0.46704 | 0.072159 | 0.69695
| Rel.Std.Dev. |  | 0.5715% | 2.233% | 0.4028% | 0.4926% | 1.48% | 1.28% |  | 0.7855% | 0.9761% | 0.2429% | 0.7391% | 0.1153% | 1.1%
| PDM3 |  | -0.5764% | 4.775% | -0.6554% | 0.9207% | -1.05% | 6.099% |  | 1.51% | 1.034% | 0.2418% | -0.2375% | -1.225% | 0.03925%
| Z-Score (Absolute) |  | 0.57 | 3.7 | 0.64 | 0.62 | 0.95 | 4.76 | NA | 1.09 | 0.71 | 0.08 | 0.3 | 1.09 | 0.08
| Analytical results for Sm2O3 in OREAS 471 (Certified Value 14.4 ppm) | [28 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 1
2
| 1 | 1 | 14.379 | 13.6869 | 15.0748 | 14.7269 | 13.5673 | 15.5154 | 14.1471 | 14.495 | 13.5673 | 14.7269 | 15.0748 | 13.6833 | 14.611 | 15.3647 | 12.5817 | 13.8513 | 16.0025 | 15.5386 | 13.7529 | 14.742 | 14.302 | 14.0312 | 13.8624 | 13.55 | 13.7574 | 14.2283 | 15.33 | 1
| 1 | 2 | 14.611 | 13.5628 | 15.0748 | 13.6833 | 14.1471 | 15.4691 | 13.9152 | 14.611 | 14.611 | 14.611 | 14.7849 | 14.7849 | 14.7849 | 13.6833 | 12.6976 | 14.5854 | 15.5386 | 13.7876 | 13.6949 | 14.8498 | 14.7026 | 13.6833 | 14.0448 | 13.9 | 13.57 | 14.6921 | 15.39 | #N/A
| 1 | 3 | 15.0748 | 14.4184 | 15.0748 | 14.2631 | 14.2631 | 15.8865 | 13.4514 | 13.7992 | 12.6396 | 14.611 | 14.9588 | 14.1471 | 14.3211 | 15.2487 | 12.1758 | 13.7792 | 15.5386 | 12.1062 | 14.3443 | 14.9913 | 15.6039 | 13.7992 | 13.4611 | 13.41 | 13.5857 | 14.6226 | 15.29 | 16
| 1 | 4 | 14.611 | 13.5088 | 12.7556 | 14.611 | 13.9152 | 15.3763 | 13.5673 | 14.2631 | 12.2918 | 14.611 | 15.5966 | 14.7269 | 14.437 | 15.4227 | 15.0168 | 13.7246 | 14.9588 | 14.1355 | 13.6833 | 14.6991 | 14.8279 | 13.7992 | 13.8 | 13.74 | 13.566 | 14.5878 | 15.36 | 14.3531
| 1 | 5 | 14.611 | 15.3152 | 13.9152 | 14.495 | 13.9152 | 15.8285 | 14.1471 | 14.379 | 15.1908 | 14.611 | 14.9009 | 14.2051 | 14.9588 | 14.9588 | 14.0312 | 14.4824 | 15.4227 | 14.3327 | 14.5182 | 14.8313 | 14.9351 | 13.6833 | 13.9557 | 13.95 | 13.5898 | 14.2979 | 15.34 | 45
| 1 | 6 | 14.2631 | 14.8899 | 13.9152 | 15.1908 | 13.7992 | 16.1532 | 14.0312 | 14.1471 | 13.6833 | 14.495 | 14.7269 | 14.1471 | 14.437 | 13.5093 | 13.3934 | 14.2645 | 14.2631 | 13.7992 | 13.8108 | 14.8243 | 16.4467 | 13.7992 | 13.5772 | 13.75 | 13.6138 | 14.7849 | 15.32
| Mean |  | 14.5916 | 14.2303 | 14.3017 | 14.495 | 13.9345 | 15.7048 | 13.8765 | 14.2824 | 13.664 | 14.611 | 15.0072 | 14.2824 | 14.5916 | 14.6979 | 13.3161 | 14.1146 | 15.2874 | 13.95 | 13.9674 | 14.823 | 15.1364 | 13.7992 | 13.7835 | 13.7167 | 13.6138 | 14.5356 | 15.3383
| Median |  | 14.611 | 14.0527 | 14.495 | 14.553 | 13.9152 | 15.672 | 13.9732 | 14.3211 | 13.6253 | 14.611 | 14.9299 | 14.1761 | 14.524 | 15.1038 | 13.0455 | 14.0579 | 15.4807 | 13.9674 | 13.7818 | 14.8278 | 14.8815 | 13.7992 | 13.8312 | 13.745 | 13.5877 | 14.6052 | 15.335
| Std Dev. |  | 0.27847 | 0.76272 | 0.94681 | 0.50279 | 0.2478 | 0.29959 | 0.29941 | 0.28796 | 1.1108 | 0.07334 | 0.3142 | 0.4127 | 0.24232 | 0.86989 | 1.0613 | 0.37772 | 0.60217 | 1.1093 | 0.36631 | 0.10103 | 0.76881 | 0.12703 | 0.22402 | 0.20569 | 0.072368 | 0.22266 | 0.034303
| Rel.Std.Dev. |  | 1.908% | 5.36% | 6.62% | 3.469% | 1.778% | 1.908% | 2.158% | 2.016% | 8.13% | 0.5019% | 2.094% | 2.89% | 1.661% | 5.918% | 7.97% | 2.676% | 3.939% | 7.952% | 2.623% | 0.6816% | 5.079% | 0.9205% | 1.625% | 1.5% | 0.5316% | 1.532% | 0.2236%
| PDM3 |  | 1.662% | -0.8557% | -0.3582% | 0.9883% | -2.917% | 9.417% | -3.321% | -0.4929% | -4.802% | 1.796% | 4.557% | -0.4929% | 1.662% | 2.402% | -7.225% | -1.662% | 6.509% | -2.809% | -2.688% | 3.273% | 5.457% | -3.859% | -3.969% | -4.434% | -5.151% | 1.271% | 6.864%
| Z-Score (Absolute) |  | 0.51 | 0.09 | 0.03 | 0.35 | 0.58 | 2.36 | 0.67 | 0 | 1.03 | 0.55 | 1.2 | 0 | 0.51 | 0.69 | 1.61 | 0.28 | 1.67 | 0.55 | 0.52 | 0.9 | 1.42 | 0.8 | 0.83 | 0.94 | 1.11 | 0.42 | 1.75
| Analytical results for Sn in OREAS 471 (Certified Value 3.17 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 27 | [6 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 3 | 3.5212 | < 20 | 3 | < 5 | < 10 | 4 | < 10 | 3 | 3.5 | 3.1 | 3.8 | 3.4 | 3.1 | 3.2578 | 3 | 3 | 2 | 2.779 | 2 | 2.9 | [6 empty] | 1
| 1 | 2 | 3 | 3.48 | < 20 | 3 | < 5 | < 10 | 4 | < 10 | 3 | 2.9 | 3.5 | 3.8 | 3.4 | 3.6 | 4.0285 | 4 | 5 | 2 | 2.897 | 2 | 3 | [6 empty] | #N/A
| 1 | 3 | 3 | 3.4146 | 90 | 3 | < 5 | < 10 | 3 | < 10 | 3 | 3.1 | 3.3 | 3.4 | 3.5 | 3.5 | 3.1875 | 4 | 3 | 3 | 3.032 | 3 | 3.2 | [6 empty] | 16
| 1 | 4 | 3 | 3.3428 | 24 | 3 | < 5 | < 10 | 5 | < 10 | 3 | 3.1 | 3.1 | 3.5 | 3.4 | 3.8 | 3.2266 | 4 | 3 | 3 | 3.61 | 4 | 2.9 | [6 empty] | 3.1671
| 1 | 5 | 3 | 3.4988 | < 20 | 3 | < 5 | < 10 | 4 | < 10 | 3 | 3.2 | 3.3 | 3.2 | 3.3 | 2.8 | 3.1096 | 4 | 4 | 2 | 2.958 | 3 | 3 | [6 empty] | 46
| 1 | 6 | 3 | 3.2557 | < 20 | 3 | < 5 | < 10 | 6 | < 10 | 3 | 3.2 | 3 | 3.3 | 3.3 | 4.4 | 3.5493 | 4 | 3 | 2 | 2.597 | 2 | 3.3
| Mean |  | 3 | 3.4188 | 57 | 3 |  |  | 4.3333 |  | 3 | 3.1667 | 3.2167 | 3.5 | 3.3833 | 3.5333 | 3.3932 | 3.8333 | 3.5 | 2.3333 | 2.9788 | 2.6667 | 3.05
| Median |  | 3 | 3.4473 | 57 | 3 |  |  | 4 |  | 3 | 3.15 | 3.2 | 3.45 | 3.4 | 3.55 | 3.2422 | 4 | 3 | 2 | 2.9275 | 2.5 | 3
| Std Dev. |  | 0 | 0.10299 | 46.669 | 0 |  |  | 1.0328 |  | 0 | 0.19664 | 0.18348 | 0.25298 | 0.075277 | 0.55737 | 0.34549 | 0.40825 | 0.83666 | 0.5164 | 0.34469 | 0.8165 | 0.16432
| Rel.Std.Dev. |  | 0% | 3.012% | 81.88% | 0% |  |  | 23.83% |  | 0% | 6.21% | 5.704% | 7.228% | 2.225% | 15.77% | 10.18% | 10.65% | 23.9% | 22.13% | 11.57% | 30.62% | 5.387%
| PDM3 |  | -5.277% | 7.947% | 1700% | -5.277% |  |  | 36.82% |  | -5.277% | -0.01491% | 1.564% | 10.51% | 6.826% | 11.56% | 7.138% | 21.03% | 10.51% | -26.33% | -5.946% | -15.8% | -3.699%
| Z-Score (Absolute) |  | 0.67 | 0.06 | 39.2 | 0.67 | 1.55 | 2.84 | 1.67 | 2.84 | 0.67 | 0.38 | 0.29 | 0.21 | 0 | 0.26 | 0.02 | 0.79 | 0.21 | 1.85 | 0.71 | 1.26 | 0.59
| Analytical results for Sr in OREAS 471 (Certified Value 445 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 26 | 27 |  |  |  |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*OES | PF*OES | PF*OES/MS |  |  |  |  | 0
0
| 1 | 1 | 436 | 461.7801 | 450 | 460 | 430 | 460 | 440 | 432 | 469 | 445.3 | 444 | 443 | 442 | 485 | 449 | 399.9897 | 454 | 444.8 | 458.5 | 452.641 | 440 | 424.4827 | 440.7 |  |  |  |  | 1
| 1 | 2 | 434 | 469.6762 | 470 | 439 | 434 | 458 | 450 | 431 | 463 | 451.4 | 430 | 449 | 429 | 459 | 453 | 415.4221 | 450 | 448.6 | 454.3 | 453.759 | 420 | 422.5387 | 445.7 |  |  |  |  | #N/A
| 1 | 3 | 430 | 449.5814 | 449 | 447 | 434 | 473 | 460 | 441 | 449 | 461.6 | 427 | 445 | 421 | 439 | 432 | 408.5803 | 453 | 431.6 | 459.2 | 467.742 | 440 | 426.8677 | 442.8 |  |  |  |  | 16
| 1 | 4 | 443 | 468.0148 | 419 | 458 | 434 | 464 | 450 | 442 | 428 | 447.8 | 436 | 443 | 435 | 459 | 448 | 394.7409 | 453 | 440.5 | 458.4 | 449.085 | 430 | 423.6547 | 440.4 |  |  |  |  | 445.1241
| 1 | 5 | 444 | 462.8627 | 436 | 456 | 430 | 468 | 450 | 443 | 462 | 448.9 | 437 | 451 | 428 | 445 | 443 | 414.4291 | 447 | 441.9 | 467.5 | 456.014 | 430 | 432.0517 | 444.6 |  |  |  |  | 47
| 1 | 6 | 432 | 466.0803 | 455 | 450 | 432 | 448 | 440 | 443 | 454 | 453.4 | 421 | 453 | 427 | 432 | 440 | 411.2426 | 452 | 441.6 | 452.8 | 433.036 | 430 | 443.2117 | 445.2
| Mean |  | 436.5 | 462.9993 | 446.5 | 451.6667 | 432.3333 | 461.8333 | 448.3333 | 438.6667 | 454.1667 | 451.4 | 432.5 | 447.3333 | 430.3333 | 453.1667 | 444.1667 | 407.4008 | 451.5 | 441.5 | 458.45 | 452.0462 | 431.6667 | 428.8012 | 443.2333
| Median |  | 435 | 464.4715 | 449.5 | 453 | 433 | 462 | 450 | 441.5 | 458 | 450.15 | 433 | 447 | 428.5 | 452 | 445.5 | 409.9115 | 452.5 | 441.75 | 458.45 | 453.2 | 430 | 425.6752 | 443.7
| Std Dev. |  | 5.7879 | 7.2218 | 17.3753 | 7.9162 | 1.9664 | 8.6814 | 7.5277 | 5.6095 | 14.6344 | 5.7352 | 8.167 | 4.274 | 7.2572 | 18.957 | 7.5211 | 8.3077 | 2.5884 | 5.6653 | 5.1306 | 11.2743 | 7.5277 | 7.8272 | 2.3001
| Rel.Std.Dev. |  | 1.326% | 1.56% | 3.891% | 1.753% | 0.4548% | 1.88% | 1.679% | 1.279% | 3.222% | 1.271% | 1.888% | 0.9554% | 1.686% | 4.183% | 1.693% | 2.039% | 0.5733% | 1.283% | 1.119% | 2.494% | 1.744% | 1.825% | 0.5189%
| PDM3 |  | -1.937% | 4.016% | 0.3091% | 1.47% | -2.874% | 3.754% | 0.721% | -1.451% | 2.031% | 1.41% | -2.836% | 0.4963% | -3.323% | 1.807% | -0.2151% | -8.475% | 1.432% | -0.8142% | 2.994% | 1.555% | -3.023% | -3.667% | -0.4248%
| Z-Score (Absolute) |  | 0.88 | 1.45 | 0 | 0.45 | 1.25 | 1.35 | 0.16 | 0.69 | 0.67 | 0.43 | 1.23 | 0.07 | 1.42 | 0.59 | 0.21 | 3.44 | 0.44 | 0.44 | 1.05 | 0.49 | 1.3 | 1.56 | 0.29
| Analytical results for Ta in OREAS 471 (Certified Value 0.88 ppm) | [28 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 27 | [5 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.9 | < 2 | < 10 | 0.9 | 1.22 | < 5 | 0.9 | 0.8 | 1.9 | 0.9 | 0.8 | 1 | 1 | 1 | 0.8 | 0.8646 | 0.7 | 2.4 | 0.8 | < 1 | 0.7 | 0.99 | [5 empty] | 1
| 1 | 2 | 0.9 | < 2 | < 10 | 0.9 | 1.23 | < 5 | 0.9 | 1 | 2 | 0.9 | 0.9 | 0.9 | 1.1 | 0.9 | 0.8 | 0.8605 | 0.8 | 2.2 | 0.8 | < 1 | 0.7 | 0.97 | [5 empty] | #N/A
| 1 | 3 | 0.9 | < 2 | < 10 | 0.9 | 0.82 | < 5 | 0.8 | 0.7 | 1.8 | 0.9 | 0.8 | 0.9 | 0.9 | 1.3 | 0.8 | 0.8654 | 0.8 | 1.3 | 0.9 | < 1 | 0.7 | 0.97 | [5 empty] | 16
| 1 | 4 | 0.9 | < 2 | < 10 | 0.9 | 1.07 | < 5 | 0.9 | 0.7 | 1.7 | 0.9 | 0.8 | 1 | 1.7 | 1 | 0.7 | 0.8497 | 0.8 | 1.7 | 0.8 | < 1 | 0.8 | 1 | [5 empty] | 0.88211
| 1 | 5 | 0.9 | < 2 | < 10 | 0.9 | 0.73 | < 5 | 1 | 1.1 | 2.1 | 0.9 | 0.9 | 1 | 1 | 0.9 | 0.8 | 0.8925 | 0.9 | 1.6 | 0.8 | < 1 | 0.7 | 1.02 | [5 empty] | 48
| 1 | 6 | 0.9 | < 2 | < 10 | 0.9 | 1.07 | < 5 | 0.9 | 0.9 | 2 | 0.9 | 0.8 | 1 | 0.9 | 0.9 | 0.8 | 0.8545 | 0.9 | 1.9 | 0.9 | < 1 | 0.7 | 1.09
| Mean |  | 0.9 |  |  | 0.9 | 1.0233 |  | 0.9 | 0.86667 | 1.9167 | 0.9 | 0.83333 | 0.96667 | 1.1 | 1 | 0.78333 | 0.86453 | 0.81667 | 1.85 | 0.83333 |  | 0.71667 | 1.0067
| Median |  | 0.9 |  |  | 0.9 | 1.07 |  | 0.9 | 0.85 | 1.95 | 0.9 | 0.8 | 1 | 1 | 0.95 | 0.8 | 0.86255 | 0.8 | 1.8 | 0.8 |  | 0.7 | 0.995
| Std Dev. |  | 0 |  |  | 0 | 0.20646 |  | 0.063246 | 0.1633 | 0.1472 | 0 | 0.05164 | 0.05164 | 0.30332 | 0.15492 | 0.040825 | 0.014966 | 0.075277 | 0.40373 | 0.05164 |  | 0.040825 | 0.045019
| Rel.Std.Dev. |  | 0% |  |  | 0% | 20.18% |  | 7.027% | 18.84% | 7.68% | 0% | 6.197% | 5.342% | 27.57% | 15.49% | 5.212% | 1.731% | 9.218% | 21.82% | 6.197% |  | 5.696% | 4.472%
| PDM3 |  | 2.029% |  |  | 2.029% | 16.01% |  | 2.029% | -1.75% | 117.3% | 2.029% | -5.529% | 9.586% | 24.7% | 13.37% | -11.2% | -1.992% | -7.419% | 109.7% | -5.529% |  | -18.76% | 14.12%
| Z-Score (Absolute) |  | 0 | 0.67 | 27.65 | 0 | 0.83 | 10.79 | 0 | 0.22 | 6.86 | 0 | 0.45 | 0.45 | 1.35 | 0.67 | 0.79 | 0.24 | 0.56 | 6.41 | 0.45 | 2.7 | 1.24 | 0.72
| Analytical results for Tb4O7 in OREAS 471 (Certified Value 1.19 ppm) | [28 empty] | Certified
Tb4O7 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*OES/MS | 2
3
| 1 | 1 | 1.1762 | < 1 | < 3.5 | 1.2232 | 1.2468 | 1.2115 | 1.1527 | 1.1762 | 1.3173 | 1.1762 | 1.188 | 1.0586 | 1.1409 | 1.2703 | 1.1409 | 1.1688 | 1.2938 | 0.96448 | 1.2115 | 1.048 | 1.587 | 1.1292 | 1.4121 | 1.24 | 1.1856 | 0.91744 | 1.11 | 1
| 1 | 2 | 1.1762 | < 1 | < 3.5 | 1.0586 | 1.1762 | 1.2703 | 1.1292 | 1.2938 | 1.3762 | 1.1762 | 1.2115 | 1.2232 | 1.1762 | 1.1762 | 1.1292 | 1.1662 | 1.2938 | 1.2703 | 1.1997 | 1.0645 | 1.4538 | 1.1292 | 1.3845 | 1.26 | 1.1889 | 0.94096 | 1.14 | #N/A
| 1 | 3 | 1.1762 | < 1 | < 3.5 | 1.1527 | 1.188 | 1.2821 | 1.1292 | 1.2938 | 1.2115 | 1.0586 | 1.1997 | 1.0821 | 1.1174 | 1.1762 | 0.99977 | 1.1516 | 1.1762 | 1.1292 | 1.235 | 1.0856 | 1.5798 | 1.1644 | 1.3861 | 1.25 | 1.1666 | 0.95272 | 1.16 | 16
| 1 | 4 | 1.1762 | < 1 | < 3.5 | 1.188 | 1.1762 | 1.2468 | 1.1644 | 1.2938 | 1.1762 | 1.1762 | 1.2115 | 1.0821 | 1.188 | 1.1997 | 1.1292 | 1.123 | 1.1762 | 1.3409 | 1.1997 | 1.0645 | 1.6973 | 1.0939 | 1.3526 | 1.25 | 1.1496 | 0.94096 | 1.25 | 1.1851
| 1 | 5 | 1.1762 | < 1 | < 3.5 | 1.1292 | 1.1644 | 1.3173 | 1.1056 | 1.2938 | 1.435 | 1.0586 | 1.188 | 1.0586 | 1.3056 | 1.1056 | 1.0115 | 1.1911 | 1.1762 | 1.1527 | 1.2703 | 1.048 | 1.5895 | 1.0939 | 1.3944 | 1.25 | 1.1651 | 0.9292 | 1.17 | 49
| 1 | 6 | 1.1762 | < 1 | < 3.5 | 1.0821 | 1.1527 | 1.2703 | 1.1292 | 1.2938 | 1.2938 | 1.1762 | 1.1292 | 1.1292 | 1.2115 | 1.1056 | 1.1174 | 1.178 | 1.1762 | 1.0233 | 1.188 | 1.0609 | 1.5812 | 1.1409 | 1.3708 | 1.22 | 1.1734 | 0.96448 | 1.19
| Mean |  | 1.1762 |  |  | 1.139 | 1.184 | 1.2664 | 1.135 | 1.2742 | 1.3017 | 1.137 | 1.188 | 1.1056 | 1.1899 | 1.1723 | 1.088 | 1.1631 | 1.2154 | 1.1468 | 1.2174 | 1.0619 | 1.5814 | 1.1252 | 1.3834 | 1.245 | 1.1716 | 0.94096 | 1.17
| Median |  | 1.1762 |  |  | 1.1409 | 1.1762 | 1.2703 | 1.1292 | 1.2938 | 1.3056 | 1.1762 | 1.1938 | 1.0821 | 1.1821 | 1.1762 | 1.1233 | 1.1675 | 1.1762 | 1.1409 | 1.2056 | 1.0627 | 1.5841 | 1.1292 | 1.3853 | 1.25 | 1.17 | 0.94096 | 1.165
| Std Dev. |  | 0 |  |  | 0.062424 | 0.03299 | 0.035416 | 0.020709 | 0.048018 | 0.097466 | 0.060739 | 0.030672 | 0.063122 | 0.065874 | 0.06209 | 0.064316 | 0.023605 | 0.060739 | 0.14266 | 0.030445 | 0.013885 | 0.077204 | 0.027501 | 0.020316 | 0.013784 | 0.014495 | 0.016634 | 0.047749
| Rel.Std.Dev. |  | 0% |  |  | 5.481% | 2.786% | 2.797% | 1.825% | 3.768% | 7.488% | 5.342% | 2.582% | 5.709% | 5.536% | 5.297% | 5.911% | 2.029% | 4.997% | 12.44% | 2.501% | 1.308% | 4.882% | 2.444% | 1.469% | 1.107% | 1.237% | 1.768% | 4.081%
| PDM3 |  | -0.7501% |  |  | -3.893% | -0.08839% | 6.859% | -4.224% | 7.521% | 9.837% | -4.058% | 0.2424% | -6.705% | 0.4079% | -1.081% | -8.194% | -1.853% | 2.558% | -3.231% | 2.724% | -10.39% | 33.44% | -5.051% | 16.73% | 5.055% | -1.142% | -20.6% | -1.273%
| Z-Score (Absolute) |  | 0.06 | 10.05 | 8.85 | 0.5 | 0.18 | 1.41 | 0.56 | 1.52 | 1.93 | 0.53 | 0.23 | 1 | 0.26 | 0 | 1.26 | 0.14 | 0.64 | 0.38 | 0.67 | 1.65 | 6.12 | 0.7 | 3.16 | 1.09 | 0.01 | 3.46 | 0.03
| Analytical results for Te in OREAS 471 (Indicative Value < 1 ppm) | [28 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 04 | 07 | 08 | 09 | 10 | 21 | 27 | [20 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | PF*OES/MS | [20 empty] | 2
2
| 1 | 1 | < 5 | < 5 | < 1 | < 1 | < 1 | 0.167 | 0.07 | [20 empty] | 1
| 1 | 2 | < 5 | < 5 | < 1 | < 1 | < 1 | 0.186 | 0.08 | [20 empty] | 2
| 1 | 3 | < 5 | < 5 | < 1 | < 1 | < 1 | 0.257 | 0.06 | [20 empty] | 16
| 1 | 4 | < 5 | < 5 | < 1 | < 1 | < 1 | 0.223 | 0.07 | [20 empty] | < 1
| 1 | 5 | < 5 | < 5 | < 1 | < 1 | < 1 | 0.181 | 0.08 | [20 empty] | 8
| 1 | 6 | < 5 | < 5 | < 1 | < 1 | < 1 | 0.139 | 0.07
| Mean | [6 empty] | 0.19217 | 0.071667
| Median | [6 empty] | 0.1835 | 0.07
| Std Dev. | [6 empty] | 0.041897 | 0.0075277
| Rel.Std.Dev. | [6 empty] | 21.8% | 10.5%
PDM3
| Z-Score (Absolute) |  | 4.38 | 4.38 | 0 | 0 | 0 | 0.67 | 0.94
| Analytical results for Th in OREAS 471 (Certified Value 22.3 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 26 | 27 |  |  |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*MS | PF*MS | PF*OES/MS |  |  |  | 1
2
| 1 | 1 | 22.8 | 22.8176 | < 50 | 22.4 | 21.7 | 26.29 | 22.4 | 22 | 21.5 | 22.7 | 22.1 | 21.5 | 23.5 | 23.4 | 21.1 | 21.6423 | 23.1 | 22.8 | 21.5 | 22.726 | 22.1 | 23.2544 | 19.6069 | 21.51 |  |  |  | 1
| 1 | 2 | 22.9 | 22.7957 | < 50 | 21.5 | 22.4 | 25.54 | 22.4 | 23 | 23 | 22.5 | 22.2 | 20.3 | 23 | 22.4 | 21 | 22.2894 | 23 | 23.6 | 21.3 | 22.604 | 21.4 | 23.2921 | 20.1189 | 21.66 |  |  |  | 15
| 1 | 3 | 23 | 22.5931 | < 50 | 22.4 | 22.5 | 27.27 | 22 | 22.3 | 20 | 22.4 | 21.8 | 21.7 | 22 | 22.1 | 20.7 | 21.473 | 22.9 | 23.9 | 21.7 | 23.115 | 21.8 | 23.0753 | 19.4852 | 21.73 |  |  |  | 16
| 1 | 4 | 22.9 | 24.2587 | < 50 | 22.1 | 21.3 | 25.23 | 22.1 | 22.3 | 19.8 | 22.5 | 22.4 | 21.1 | 23.7 | 22.4 | 21.3 | 21.4691 | 22.8 | 23.2 | 21.1 | 22.75 | 21.6 | 23.0464 | 19.1685 | 21.75 |  |  |  | 22.2751
| 1 | 5 | 23.2 | 22.8116 | < 50 | 21.9 | 21 | 25.9 | 22.4 | 22.3 | 23.6 | 22.6 | 22.6 | 22.1 | 23.8 | 22.1 | 21.2 | 22.4879 | 23.2 | 22.8 | 22.2 | 22.533 | 21.8 | 23.2599 | 19.8043 | 22.49 |  |  |  | 50
| 1 | 6 | 22.7 | 22.3847 | < 50 | 22.4 | 22.3 | 25.15 | 22.4 | 22.5 | 21.2 | 22.4 | 21.8 | 21.7 | 22.9 | 21.5 | 20.9 | 21.9294 | 22.8 | 22.2 | 21.6 | 22.646 | 22.3 | 23.3317 | 19.656 | 21.76
| Mean |  | 22.9167 | 22.9436 |  | 22.1167 | 21.8667 | 25.8967 | 22.2833 | 22.4 | 21.5167 | 22.5167 | 22.15 | 21.4 | 23.15 | 22.3167 | 21.0333 | 21.8818 | 22.9667 | 23.0833 | 21.5667 | 22.729 | 21.8333 | 23.21 | 19.64 | 21.8167
| Median |  | 22.9 | 22.8036 |  | 22.25 | 22 | 25.72 | 22.4 | 22.3 | 21.35 | 22.5 | 22.15 | 21.6 | 23.25 | 22.25 | 21.05 | 21.7858 | 22.95 | 23 | 21.55 | 22.686 | 21.8 | 23.2572 | 19.6315 | 21.74
| Std Dev. |  | 0.17224 | 0.66631 |  | 0.3656 | 0.62823 | 0.79623 | 0.18348 | 0.33466 | 1.542 | 0.1169 | 0.32094 | 0.62929 | 0.67157 | 0.62423 | 0.21602 | 0.43141 | 0.1633 | 0.61455 | 0.37771 | 0.20512 | 0.3266 | 0.11909 | 0.31748 | 0.34268
| Rel.Std.Dev. |  | 0.7516% | 2.904% |  | 1.653% | 2.873% | 3.075% | 0.8234% | 1.494% | 7.166% | 0.5192% | 1.449% | 2.941% | 2.901% | 2.797% | 1.027% | 1.972% | 0.711% | 2.662% | 1.751% | 0.9025% | 1.496% | 0.5131% | 1.616% | 1.571%
| PDM3 |  | 2.88% | 3.001% |  | -0.7115% | -1.834% | 16.26% | 0.03676% | 0.5605% | -3.405% | 1.084% | -0.5618% | -3.929% | 3.927% | 0.1864% | -5.575% | -1.766% | 3.104% | 3.628% | -3.181% | 2.037% | -1.983% | 4.197% | -11.83% | -2.058%
| Z-Score (Absolute) |  | 0.66 | 0.69 | 2.89 | 0.2 | 0.46 | 3.85 | 0.02 | 0.11 | 0.84 | 0.23 | 0.16 | 0.96 | 0.91 | 0.02 | 1.36 | 0.45 | 0.71 | 0.84 | 0.78 | 0.46 | 0.5 | 0.97 | 2.85 | 0.52
| Analytical results for Ti in OREAS 471 (Certified Value 0.566 wt.%) | [28 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 23 | 24 | 26 | 27 | [6 empty] | wt.%
|  |  | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.5509 | 0.6176 | 0.58 | 0.5549 | 0.59 | 0.56 | 0.57 | 0.54 | 0.576 | 0.62 | 0.6 | 0.54 | 0.6 | 0.58 | 0.52982 | 0.55 | 0.57 | 0.568 | 0.51467 | 0.57982 | 0.528 | [6 empty] | 1
| 1 | 2 | 0.5521 | 0.6628 | 0.57 | 0.5636 | 0.58 | 0.54 | 0.57 | 0.54 | 0.581 | 0.61 | 0.54 | 0.53 | 0.57 | 0.59 | 0.54253 | 0.56 | 0.57 | 0.552 | 0.51924 | 0.57618 | 0.534 | [6 empty] | #N/A
| 1 | 3 | 0.5521 | 0.6253 | 0.57 | 0.55 | 0.58 | 0.55 | 0.57 | 0.54 | 0.581 | 0.6 | 0.61 | 0.52 | 0.56 | 0.57 | 0.53246 | 0.57 | 0.57 | 0.569 | 0.5143 | 0.58659 | 0.546 | [6 empty] | 16
| 1 | 4 | 0.5599 | 0.5702 | 0.57 | 0.5583 | 0.57 | 0.56 | 0.59 | 0.5 | 0.576 | 0.62 | 0.59 | 0.55 | 0.58 | 0.58 | 0.51879 | 0.56 | 0.57 | 0.567 | 0.51528 | 0.58706 | 0.528 | [6 empty] | 0.56583
| 1 | 5 | 0.5533 | 0.6189 | 0.58 | 0.5611 | 0.59 | 0.55 | 0.58 | 0.55 | 0.587 | 0.61 | 0.6 | 0.54 | 0.55 | 0.58 | 0.54727 | 0.57 | 0.57 | 0.567 | 0.51613 | 0.59302 | 0.54 | [6 empty] | 51
| 1 | 6 | 0.5515 | 0.6446 | 0.56 | 0.5719 | 0.55 | 0.55 | 0.6 | 0.53 | 0.581 | 0.59 | 0.6 | 0.52 | 0.54 | 0.58 | 0.54277 | 0.56 | 0.564 | 0.59 | 0.51551 | 0.58691 | 0.564
| Mean |  | 0.5533 | 0.62323 | 0.57167 | 0.55997 | 0.57667 | 0.55167 | 0.58 | 0.53333 | 0.58033 | 0.60833 | 0.59 | 0.53333 | 0.56667 | 0.58 | 0.53561 | 0.56167 | 0.569 | 0.56883 | 0.51585 | 0.58493 | 0.54
| Median |  | 0.5521 | 0.6221 | 0.57 | 0.5597 | 0.58 | 0.55 | 0.575 | 0.54 | 0.581 | 0.61 | 0.6 | 0.535 | 0.565 | 0.58 | 0.5375 | 0.56 | 0.57 | 0.5675 | 0.5154 | 0.58675 | 0.537
| Std Dev. |  | 0.0033299 | 0.031271 | 0.0075277 | 0.0075455 | 0.015055 | 0.0075277 | 0.012649 | 0.017512 | 0.0040825 | 0.01169 | 0.025298 | 0.012111 | 0.021602 | 0.0063246 | 0.010596 | 0.0075277 | 0.0024495 | 0.012156 | 0.001778 | 0.0059899 | 0.013682
| Rel.Std.Dev. |  | 0.6018% | 5.018% | 1.317% | 1.347% | 2.611% | 1.365% | 2.181% | 3.283% | 0.7035% | 1.922% | 4.288% | 2.271% | 3.812% | 1.09% | 1.978% | 1.34% | 0.4305% | 2.137% | 0.3447% | 1.024% | 2.534%
| PDM3 |  | -2.214% | 10.15% | 1.032% | -1.036% | 1.916% | -2.503% | 2.505% | -5.743% | 2.564% | 7.512% | 4.272% | -5.743% | 0.1482% | 2.505% | -5.341% | -0.7354% | 0.5606% | 0.5311% | -8.832% | 3.376% | -4.565%
| Z-Score (Absolute) |  | 0.67 | 2.43 | 0.14 | 0.38 | 0.36 | 0.75 | 0.51 | 1.56 | 0.53 | 1.77 | 0.95 | 1.56 | 0.08 | 0.51 | 1.46 | 0.3 | 0 | 0.01 | 2.34 | 0.73 | 1.29
| Analytical results for Tl in OREAS 471 (Certified Value 0.64 ppm) | [28 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 04 | 06 | 08 | 09 | 10 | 18 | 21 | 24 | 26 | 27 | [16 empty] | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | PF*MS | PF*OES/MS | [16 empty] | 2
2
| 1 | 1 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.6 | 0.7 | 0.62 | 0.46775 | 0.49103 | 0.56 | [16 empty] | 1
| 1 | 2 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.8 | 0.617 | 0.47603 | 0.53904 | 0.56 | [16 empty] | 16
| 1 | 3 | 0.7 | 0.7 | 0.7 | 0.7 | 0.6 | 0.5 | 0.8 | 0.644 | 0.49397 | 0.48047 | 0.56 | [16 empty] | 16
| 1 | 4 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.6 | 0.8 | 0.635 | 0.46893 | 0.48014 | 0.55 | [16 empty] | 0.64193
| 1 | 5 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.8 | 0.659 | 0.53277 | 0.48989 | 0.55 | [16 empty] | 52
| 1 | 6 | 0.7 | 0.8 | 0.7 | 0.7 | 0.8 | 0.6 | 0.7 | 0.58 | 0.45232 | 0.50022 | 0.56
| Mean |  | 0.7 | 0.71667 | 0.7 | 0.7 | 0.7 | 0.61667 | 0.76667 | 0.62583 | 0.48196 | 0.4968 | 0.55667
| Median |  | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.6 | 0.8 | 0.6275 | 0.47248 | 0.49046 | 0.56
| Std Dev. |  | 1.2162e-16 | 0.040825 | 1.2162e-16 | 1.2162e-16 | 0.063246 | 0.075277 | 0.05164 | 0.027301 | 0.028323 | 0.022006 | 0.005164
| Rel.Std.Dev. |  | 1.737e-14% | 5.696% | 1.737e-14% | 1.737e-14% | 9.035% | 12.21% | 6.736% | 4.362% | 5.877% | 4.429% | 0.9277%
| PDM3 |  | 9.046% | 11.64% | 9.046% | 9.046% | 9.046% | -3.936% | 19.43% | -2.508% | -24.92% | -22.61% | -13.28%
| Z-Score (Absolute) |  | 0 | 0.17 | 0 | 0 | 0 | 0.84 | 0.67 | 0.75 | 2.21 | 2.06 | 1.45
| Analytical results for Tm2O3 in OREAS 471 (Certified Value 0.39 ppm) | [28 empty] | Certified
Tm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.37689 | 0.40572 | < 2.3 | 0.38831 | 0.35405 | 0.39973 | 0.434 | 0.34263 | 0.38831 | 0.34263 | 0.39973 | 0.41116 | 0.41116 | 0.35405 | 0.36547 | 0.39699 | 0.45684 | 0.39973 | 0.37689 | 0.426 | 0.3253 | 0.39973 | 0.37732 | 0.434 | 0.36109 | 0.36547 |  | 1
| 1 | 2 | 0.38831 | 0.39701 | < 2.3 | 0.36547 | 0.42258 | 0.38831 | 0.37689 | 0.45684 | 0.37689 | 0.34263 | 0.42258 | 0.39973 | 0.39973 | 0.39973 | 0.44542 | 0.39323 | 0.34263 | 0.38831 | 0.37689 | 0.42372 | 0.30754 | 0.37689 | 0.38642 | 0.434 | 0.36249 | 0.36547 |  | #N/A
| 1 | 3 | 0.38831 | 0.37358 | < 2.3 | 0.37689 | 0.38831 | 0.42258 | 0.37689 | 0.45684 | 0.34263 | 0.34263 | 0.39973 | 0.37689 | 0.39973 | 0.37689 | 0.37689 | 0.36993 | 0.45684 | 0.39973 | 0.37689 | 0.42829 | 0.34149 | 0.37689 | 0.36836 | 0.44542 | 0.34835 | 0.37689 |  | 16
| 1 | 4 | 0.38831 | 0.386 | < 2.3 | 0.36547 | 0.37689 | 0.39973 | 0.37689 | 0.34263 | 0.34263 | 0.45684 | 0.39973 | 0.38831 | 0.41116 | 0.38831 | 0.36547 | 0.38489 | 0.34263 | 0.42258 | 0.37689 | 0.426 | 0.34824 | 0.37689 | 0.39258 | 0.45684 | 0.34791 | 0.37689 |  | 0.39002
| 1 | 5 | 0.37689 | 0.40674 | < 2.3 | 0.37689 | 0.38831 | 0.41116 | 0.38831 | 0.45684 | 0.41116 | 0.45684 | 0.42258 | 0.34263 | 0.434 | 0.36547 | 0.35405 | 0.41207 | 0.45684 | 0.47968 | 0.39973 | 0.43057 | 0.34321 | 0.37689 | 0.39562 | 0.41116 | 0.36038 | 0.36547 |  | 53
| 1 | 6 | 0.39973 | 0.41576 | < 2.3 | 0.36547 | 0.37689 | 0.434 | 0.37689 | 0.45684 | 0.37689 | 0.34263 | 0.37689 | 0.37689 | 0.36547 | 0.34263 | 0.45684 | 0.39928 | 0.45684 | 0.39973 | 0.37689 | 0.43628 | 0.33869 | 0.39973 | 0.38055 | 0.38831 | 0.34621 | 0.37689
| Mean |  | 0.38641 | 0.39747 |  | 0.37309 | 0.38451 | 0.40925 | 0.38831 | 0.41877 | 0.37309 | 0.3807 | 0.40354 | 0.3826 | 0.40354 | 0.37118 | 0.39402 | 0.39273 | 0.41877 | 0.41496 | 0.3807 | 0.42848 | 0.33408 | 0.38451 | 0.38348 | 0.42829 | 0.3544 | 0.37118
| Median |  | 0.38831 | 0.40136 |  | 0.37118 | 0.3826 | 0.40545 | 0.37689 | 0.45684 | 0.37689 | 0.34263 | 0.39973 | 0.3826 | 0.40545 | 0.37118 | 0.37118 | 0.39511 | 0.45684 | 0.39973 | 0.37689 | 0.42715 | 0.34009 | 0.37689 | 0.38349 | 0.434 | 0.35436 | 0.37118
| Std Dev. |  | 0.0085974 | 0.015433 |  | 0.0093252 | 0.022458 | 0.016811 | 0.022842 | 0.058978 | 0.026703 | 0.058978 | 0.017195 | 0.023683 | 0.022458 | 0.021367 | 0.044965 | 0.014261 | 0.058978 | 0.033623 | 0.0093252 | 0.0044771 | 0.015109 | 0.011796 | 0.010139 | 0.02476 | 0.0076393 | 0.0062555
| Rel.Std.Dev. |  | 2.225% | 3.883% |  | 2.499% | 5.841% | 4.108% | 5.882% | 14.08% | 7.157% | 15.49% | 4.261% | 6.19% | 5.565% | 5.756% | 11.41% | 3.631% | 14.08% | 8.103% | 2.449% | 1.045% | 4.523% | 3.068% | 2.644% | 5.781% | 2.156% | 1.685%
| PDM3 |  | -0.9247% | 1.91% |  | -4.341% | -1.413% | 4.932% | -0.4367% | 7.372% | -4.341% | -2.389% | 3.468% | -1.901% | 3.468% | -4.829% | 1.027% | 0.6956% | 7.372% | 6.396% | -2.389% | 9.861% | -14.34% | -1.413% | -1.677% | 9.812% | -9.131% | -4.829%
| Z-Score (Absolute) |  | 0 | 0.56 | NA | 0.67 | 0.1 | 1.16 | 0.1 | 1.64 | 0.67 | 0.29 | 0.87 | 0.19 | 0.87 | 0.77 | 0.39 | 0.32 | 1.64 | 1.44 | 0.29 | 2.13 | 2.65 | 0.1 | 0.15 | 2.12 | 1.62 | 0.77
Results from laboratory 1.03 were removed due to their systematic bias.
| Analytical results for TREO in OREAS 471 (Certified Value 537 ppm) | [28 empty] | Certified
TREO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 535.9592 | 543.2643 | 513.564 | 571.28 | 526.5104 | 530.8161 | 548.6793 | 539.1419 | 554.6581 | 523.529 | 558.2194 | 561.774 | 521.7933 | 525.5877 | 561.0934 | 540.3292 | 531.7518 | 540.0326 | 589.6563 | 535.7652 | 564.9723 | 548.6869 | 540.6842 | 536.2275 |  |  |  | 1
| 1 | 2 | 532.3827 | 521.4681 | 528.4481 | 563.5348 | 538.9879 | 545.9709 | 582.3499 | 541.3345 | 549.8327 | 530.4944 | 545.3978 | 538.023 | 526.4805 | 543.5973 | 557.9789 | 555.0999 | 530.9542 | 541.0905 | 598.5991 | 522.0614 | 562.8695 | 549.6305 | 533.9846 | 542.9391 |  |  |  | #N/A
| 1 | 3 | 544.3504 | 535.1575 | 524.2034 | 569.414 | 533.3802 | 538.2631 | 530.0005 | 543.4827 | 538.5185 | 529.7485 | 519.5442 | 531.1149 | 519.8663 | 519.4297 | 562.6874 | 531.5403 | 545.7956 | 547.4557 | 587.7368 | 537.7091 | 567.478 | 548.4883 | 534.6493 | 542.2006 |  |  |  | 16
| 1 | 4 | 544.0481 | 532.2587 | 518.1171 | 557.1884 | 531.9646 | 539.3989 | 507.7667 | 544.8345 | 549.1588 | 527.0016 | 547.5699 | 543.88 | 539.5855 | 519.239 | 550.566 | 531.065 | 529.6462 | 540.4159 | 615.2521 | 522.2928 | 564.8016 | 549.4965 | 526.888 | 537.4811 |  |  |  | 537.2279
| 1 | 5 | 545.2524 | 536.234 | 525.2109 | 566.3029 | 536.4243 | 547.1226 | 611.018 | 544.3099 | 552.029 | 534.1817 | 536.5993 | 524.6682 | 530.547 | 545.147 | 556.865 | 539.0428 | 560.5296 | 537.3702 | 604.2823 | 524.0921 | 566.6427 | 548.638 | 532.5816 | 536.4481 |  |  |  | 54
| 1 | 6 | 531.7112 | 549.4137 | 524.115 | 557.9879 | 527.8261 | 539.3576 | 547.7537 | 540.0436 | 530.0714 | 530.9733 | 531.1862 | 514.1753 | 534.4302 | 530.5826 | 545.2606 | 527.7083 | 535.7427 | 539.2995 | 601.0094 | 538.6563 | 566.5061 | 548.7433 | 532.9335 | 545.857
| Mean |  | 538.9507 | 536.2994 | 522.2764 | 564.2847 | 532.5156 | 540.1549 | 554.5947 | 542.1912 | 545.7114 | 529.3214 | 539.7528 | 535.6059 | 528.7838 | 530.5972 | 555.7419 | 537.4643 | 539.07 | 540.944 | 599.4227 | 530.0962 | 565.545 | 548.9472 | 533.6202 | 540.1923
| Median |  | 540.0037 | 535.6957 | 524.1592 | 564.9189 | 532.6724 | 539.3782 | 548.2165 | 542.4086 | 549.4957 | 530.1215 | 540.9986 | 534.5689 | 528.5138 | 528.0851 | 557.422 | 535.2915 | 533.7473 | 540.2242 | 599.8043 | 529.9287 | 565.7392 | 548.7151 | 533.4591 | 539.8409
| Std Dev. |  | 6.3143 | 9.5718 | 5.4222 | 5.829 | 4.8223 | 5.9055 | 36.9688 | 2.3575 | 9.4358 | 3.6563 | 13.6067 | 16.4811 | 7.5546 | 11.4849 | 6.6301 | 9.9291 | 12.042 | 3.4354 | 10.0928 | 8.0607 | 1.6676 | 0.48665 | 4.4264 | 4.0188
| Rel.Std.Dev. |  | 1.172% | 1.785% | 1.038% | 1.033% | 0.9056% | 1.093% | 6.666% | 0.4348% | 1.729% | 0.6907% | 2.521% | 3.077% | 1.429% | 2.165% | 1.193% | 1.847% | 2.234% | 0.6351% | 1.684% | 1.521% | 0.2949% | 0.08865% | 0.8295% | 0.744%
| PDM3 |  | 0.3207% | -0.1728% | -2.783% | 5.036% | -0.8772% | 0.5448% | 3.233% | 0.9239% | 1.579% | -1.472% | 0.47% | -0.3019% | -1.572% | -1.234% | 3.446% | 0.04399% | 0.3429% | 0.6917% | 11.58% | -1.328% | 5.271% | 2.181% | -0.6715% | 0.5518%
| Z-Score (Absolute) |  | 0.05 | 0.32 | 1.75 | 2.54 | 0.7 | 0.08 | 1.55 | 0.28 | 0.64 | 1.03 | 0.03 | 0.39 | 1.09 | 0.9 | 1.67 | 0.2 | 0.03 | 0.16 | 6.13 | 0.95 | 2.67 | 0.97 | 0.59 | 0.08
| Analytical results for TREO+Y2O3 in OREAS 471 (Certified Value 571 ppm) | [28 empty] | Certified
TREO+Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | BF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 566.4368 | 575.2658 | 542.1748 | 604.2974 | 556.734 | 562.6906 | 579.2839 | 571.2704 | 588.0565 | 554.3876 | 590.3478 | 589.874 | 552.2709 | 555.6383 | 593.0949 | 574.2355 | 562.1024 | 571.6379 | 626.6392 | 558.8774 | 599.2732 | 580.8569 | 575.1761 | 568.356 |  |  |  | 1
| 1 | 2 | 562.8603 | 552.1997 | 558.5956 | 597.8221 | 568.5766 | 578.0994 | 613.5894 | 572.701 | 584.3739 | 561.353 | 577.1453 | 563.823 | 558.99 | 574.051 | 589.4724 | 588.2443 | 560.5429 | 573.2431 | 635.7989 | 543.1417 | 597.4955 | 582.4805 | 567.6397 | 574.6866 |  |  |  | #N/A
| 1 | 3 | 573.5581 | 566.397 | 553.1444 | 601.1615 | 562.2069 | 569.7567 | 564.9228 | 574.8493 | 571.6629 | 561.496 | 549.3868 | 556.2149 | 550.4709 | 549.5183 | 595.1969 | 564.4307 | 576.5272 | 579.654 | 622.3943 | 567.0438 | 602.2502 | 581.0583 | 569.8488 | 574.0751 |  |  |  | 16
| 1 | 4 | 573.2558 | 564.1332 | 546.8676 | 590.2058 | 560.6643 | 571.7814 | 536.0855 | 576.836 | 582.5571 | 559.5111 | 579.8253 | 569.28 | 570.3171 | 548.9294 | 583.4564 | 564.2094 | 560.1238 | 572.0872 | 652.9088 | 554.1673 | 598.8767 | 581.9865 | 562.0017 | 569.1017 |  |  |  | 571.3211
| 1 | 5 | 575.73 | 567.3466 | 554.4821 | 598.0504 | 565.378 | 579.886 | 642.6386 | 576.8194 | 585.6813 | 564.9132 | 567.2039 | 548.9682 | 562.1675 | 575.9094 | 588.8665 | 573.0761 | 591.5152 | 568.8345 | 640.0448 | 556.0936 | 600.8436 | 580.948 | 566.9153 | 567.9417 |  |  |  | 55
| 1 | 6 | 559.649 | 581.2882 | 554.1481 | 589.7354 | 556.3988 | 571.867 | 578.3583 | 570.7752 | 562.5809 | 562.8478 | 562.1718 | 539.0753 | 568.0825 | 561.1987 | 576.8811 | 561.4877 | 565.9664 | 570.7066 | 640.4902 | 564.5623 | 600.9575 | 580.8033 | 567.9575 | 578.2394
| Mean |  | 568.5817 | 567.7717 | 551.5688 | 596.8788 | 561.6598 | 572.3468 | 585.8131 | 573.8752 | 579.1521 | 560.7515 | 571.0135 | 561.2059 | 560.3832 | 560.8742 | 587.8281 | 570.9473 | 569.463 | 572.6939 | 636.3794 | 557.3144 | 599.9495 | 581.3556 | 568.2565 | 572.0667
| Median |  | 569.8463 | 566.8718 | 553.6463 | 597.9363 | 561.4356 | 571.8242 | 578.8211 | 573.7751 | 583.4655 | 561.4245 | 572.1746 | 560.0189 | 560.5788 | 558.4185 | 589.1695 | 568.7534 | 564.0344 | 571.8625 | 637.9219 | 557.4855 | 600.0584 | 581.0031 | 567.7986 | 571.5884
| Std Dev. |  | 6.555 | 9.9665 | 5.9556 | 5.8514 | 4.7937 | 6.1714 | 37.4382 | 2.6891 | 9.9144 | 3.5963 | 14.4877 | 17.6541 | 8.0889 | 11.8224 | 6.7048 | 9.9185 | 12.4048 | 3.7168 | 10.9027 | 8.5076 | 1.7169 | 0.70309 | 4.285 | 4.209
| Rel.Std.Dev. |  | 1.153% | 1.755% | 1.08% | 0.9803% | 0.8535% | 1.078% | 6.391% | 0.4686% | 1.712% | 0.6413% | 2.537% | 3.146% | 1.443% | 2.108% | 1.141% | 1.737% | 2.178% | 0.649% | 1.713% | 1.527% | 0.2862% | 0.1209% | 0.7541% | 0.7357%
| PDM3 |  | -0.4795% | -0.6213% | -3.457% | 4.473% | -1.691% | 0.1795% | 2.537% | 0.447% | 1.371% | -1.85% | -0.05384% | -1.77% | -1.915% | -1.829% | 2.889% | -0.06543% | -0.3252% | 0.2403% | 11.39% | -2.452% | 5.011% | 1.756% | -0.5364% | 0.1305%
| Z-Score (Absolute) |  | 0.17 | 0.23 | 1.37 | 1.83 | 0.66 | 0.1 | 1.05 | 0.2 | 0.58 | 0.72 | 0 | 0.69 | 0.75 | 0.71 | 1.19 | 0 | 0.11 | 0.12 | 4.62 | 0.97 | 2.05 | 0.73 | 0.19 | 0.08
| Analytical results for U in OREAS 471 (Certified Value 2.61 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 27 |  |  |  |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.6 | < 5 | < 50 | 2.7 | 2.57 | 2.87 | 2.6 | 2.6 | 3 | 2.6 | 2.46 | 2.52 | 2.61 | 2.74 | 2.47 | 2.661 | 2.7 | 2.4 | 2.5 | 2.805 | 2.64 | 2.1346 | 2.5 |  |  |  |  | 1
| 1 | 2 | 2.6 | < 5 | < 50 | 2.6 | 2.62 | 2.95 | 2.7 | 2.7 | 3 | 2.6 | 2.54 | 2.44 | 2.66 | 2.54 | 2.11 | 2.6633 | 2.7 | 2.6 | 2.5 | 2.79 | 2.69 | 2.1976 | 2.6 |  |  |  |  | #N/A
| 1 | 3 | 2.7 | < 5 | < 50 | 2.5 | 2.56 | 2.87 | 2.7 | 2.6 | 3 | 2.5 | 2.48 | 2.5 | 2.47 | 2.89 | 2.44 | 2.768 | 2.7 | 2.4 | 2.5 | 2.891 | 2.6 | 2.127 | 2.6 |  |  |  |  | 16
| 1 | 4 | 2.6 | < 5 | < 50 | 2.8 | 2.6 | 2.91 | 2.6 | 2.6 | 3 | 2.6 | 2.52 | 2.44 | 2.63 | 2.65 | 2.31 | 2.6092 | 2.7 | 2.7 | 2.5 | 2.831 | 2.62 | 2.2065 | 2.5 |  |  |  |  | 2.6086
| 1 | 5 | 2.7 | < 5 | < 50 | 2.6 | 2.58 | 2.82 | 2.7 | 2.7 | 3 | 2.7 | 2.56 | 2.5 | 2.72 | 2.49 | 2.36 | 2.526 | 2.8 | 2.7 | 2.6 | 2.811 | 2.63 | 2.1678 | 2.5 |  |  |  |  | 56
| 1 | 6 | 2.6 | < 5 | < 50 | 2.7 | 2.56 | 2.81 | 2.7 | 2.8 | 3 | 2.6 | 2.38 | 2.46 | 2.71 | 2.44 | 2.36 | 2.7085 | 2.8 | 2.7 | 2.5 | 2.816 | 2.67 | 2.1252 | 2.8
| Mean |  | 2.6333 |  |  | 2.65 | 2.5817 | 2.8717 | 2.6667 | 2.6667 | 3 | 2.6 | 2.49 | 2.4767 | 2.6333 | 2.625 | 2.3417 | 2.656 | 2.7333 | 2.5833 | 2.5167 | 2.824 | 2.6417 | 2.1598 | 2.5833
| Median |  | 2.6 |  |  | 2.65 | 2.575 | 2.87 | 2.7 | 2.65 | 3 | 2.6 | 2.5 | 2.48 | 2.645 | 2.595 | 2.36 | 2.6622 | 2.7 | 2.65 | 2.5 | 2.8135 | 2.635 | 2.1512 | 2.55
| Std Dev. |  | 0.05164 |  |  | 0.10488 | 0.024014 | 0.053072 | 0.05164 | 0.08165 | 0 | 0.063246 | 0.065422 | 0.034448 | 0.090921 | 0.16956 | 0.12766 | 0.083029 | 0.05164 | 0.1472 | 0.040825 | 0.035463 | 0.033116 | 0.036291 | 0.1169
| Rel.Std.Dev. |  | 1.961% |  |  | 3.958% | 0.9302% | 1.848% | 1.936% | 3.062% | 0% | 2.433% | 2.627% | 1.391% | 3.453% | 6.459% | 5.452% | 3.126% | 1.889% | 5.698% | 1.622% | 1.256% | 1.254% | 1.68% | 4.525%
| PDM3 |  | 0.948% |  |  | 1.587% | -1.033% | 10.08% | 2.226% | 2.226% | 15% | -0.3298% | -4.547% | -5.058% | 0.948% | 0.6285% | -10.23% | 1.817% | 4.781% | -0.9687% | -3.524% | 8.257% | 1.267% | -17.21% | -0.9687%
| Z-Score (Absolute) |  | 0.06 | 1.87 | 323.25 | 0.3 | 0.69 | 3.5 | 0.54 | 0.54 | NA | 0.42 | 2.01 | 2.2 | 0.06 | 0.06 | 4.15 | 0.39 | 1.51 | 0.66 | 1.63 | 2.82 | 0.18 | 6.78 | 0.66
Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 471 (Certified Value 104 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 26 | 27 |  |  |  |  | ppm
|  |  | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 0
0
| 1 | 1 | 106 | 101 | 108 | 96 | 109 | 100 | 128 | 95 | 104 | 126 | 109 | 102 | 124 | 111 | 96.5719 | < 50 | 100 | 113 | 103.844 | 71 | 100.5645 | 100.205 | 105 |  |  |  |  | 1
| 1 | 2 | 107 | 105 | 106 | 100 | 111 | 100 | 132 | 100 | 106 | 123 | 108 | 99 | 114 | 112 | 101.8335 | 68 | 102 | 114 | 111.486 | 65 | 99.766 | 100.907 | 106 |  |  |  |  | #N/A
| 1 | 3 | 106 | 97 | 108 | 98 | 114 | 100 | 112 | 95 | 107 | 122 | 113 | 98 | 111 | 107 | 99.4768 | 75 | 97 | 113 | 101.223 | 93 | 100.0415 | 101.978 | 106 |  |  |  |  | 16
| 1 | 4 | 105 | 92 | 107 | 95 | 105 | 100 | 126 | 90 | 105 | 123 | 108 | 102 | 114 | 112 | 96.972 | 76 | 100 | 110 | 107.712 | 104 | 99.766 | 99.656 | 105 |  |  |  |  | 103.9331
| 1 | 5 | 107 | 95 | 108 | 98 | 107 | 100 | 128 | 99 | 108 | 121 | 112 | 98 | 110 | 110 | 100.9543 | 88 | 100 | 115 | 99.995 | 103 | 100.5355 | 103.139 | 105 |  |  |  |  | 57
| 1 | 6 | 107 | 100 | 106 | 99 | 104 | 100 | 126 | 95 | 106 | 122 | 110 | 98 | 108 | 112 | 100.598 | 74 | 99 | 123 | 99.907 | 81 | 101.367 | 104.471 | 106
| Mean |  | 106.3333 | 98.3333 | 107.1667 | 97.6667 | 108.3333 | 100 | 125.3333 | 95.6667 | 106 | 122.8333 | 110 | 99.5 | 113.5 | 110.6667 | 99.4011 | 76.2 | 99.6667 | 114.6667 | 104.0278 | 86.1667 | 100.3401 | 101.726 | 105.5
| Median |  | 106.5 | 98.5 | 107.5 | 98 | 108 | 100 | 127 | 95 | 106 | 122.5 | 109.5 | 98.5 | 112.5 | 111.5 | 100.0374 | 75 | 100 | 113.5 | 102.5335 | 87 | 100.2885 | 101.4425 | 105.5
| Std Dev. |  | 0.8165 | 4.6332 | 0.98319 | 1.8619 | 3.7771 | 0 | 6.8896 | 3.559 | 1.4142 | 1.7224 | 2.0976 | 1.9748 | 5.648 | 1.9664 | 2.1759 | 7.2938 | 1.633 | 4.4121 | 4.6977 | 16.4489 | 0.61505 | 1.8365 | 0.54772
| Rel.Std.Dev. |  | 0.7679% | 4.712% | 0.9174% | 1.906% | 3.487% | 0% | 5.497% | 3.72% | 1.334% | 1.402% | 1.907% | 1.985% | 4.976% | 1.777% | 2.189% | 9.572% | 1.638% | 3.848% | 4.516% | 19.09% | 0.613% | 1.805% | 0.5192%
| PDM3 |  | 2.309% | -5.388% | 3.111% | -6.029% | 4.234% | -3.784% | 20.59% | -7.954% | 1.989% | 18.19% | 5.837% | -4.265% | 9.205% | 6.479% | -4.36% | -26.68% | -4.105% | 10.33% | 0.09118% | -17.09% | -3.457% | -2.124% | 1.508%
| Z-Score (Absolute) |  | 0.34 | 0.83 | 0.46 | 0.93 | 0.63 | 0.59 | 3.11 | 1.22 | 0.29 | 2.74 | 0.87 | 0.66 | 1.38 | 0.97 | 0.67 | 5.3 | 0.64 | 1.55 | 0 | 2.6 | 0.54 | 0.34 | 0.21
| Analytical results for W in OREAS 471 (Certified Value 4.29 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 27 | [6 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 4 | < 10 | < 20 | 4.2 | 2.3 | < 10 | 5 | 6 | 4 | 4.6 | 4.9 | 4.4 | 4.5 | 4.4 | 3.6906 | 2 | 4.2 | 4.3 | 4.187 | 5 | 8.2 | [6 empty] | 1
| 1 | 2 | 4 | < 10 | < 20 | 3.9 | 1.5 | < 10 | 5 | 5 | 4 | 4.4 | 4.4 | 4.4 | 4.4 | 4.6 | 3.7033 | 3 | 3.7 | 4.3 | 3.945 | 5 | 7.8 | [6 empty] | #N/A
| 1 | 3 | 4 | < 10 | 119 | 4.1 | 1.3 | < 10 | 4 | 3 | 4 | 4.5 | 4.4 | 4.4 | 4.8 | 4.4 | 3.8384 | 3 | 4.4 | 4.3 | 4.251 | 5 | 7.9 | [6 empty] | 16
| 1 | 4 | 4 | < 10 | < 20 | 4 | 1.5 | < 10 | 4 | 4 | 4 | 4.8 | 4.8 | 4 | 4.4 | 4.1 | 3.5503 | 3 | 4.7 | 4.2 | 3.938 | 5 | 8 | [6 empty] | 4.2914
| 1 | 5 | 4 | < 10 | < 20 | 4.1 | 1 | < 10 | 5 | 5 | 4 | 4.6 | 4.9 | 4.7 | 4 | 4.1 | 4.4724 | 3 | 4.8 | 4.5 | 4.038 | 5 | 8.5 | [6 empty] | 58
| 1 | 6 | 4 | < 10 | < 20 | 3.9 | 1.3 | < 10 | 5 | 6 | 4 | 4.5 | 4.6 | 4.1 | 4.3 | 4.4 | 3.7336 | 3 | 5.4 | 4.3 | 4.076 | 5 | 8.3
| Mean |  | 4 |  | 119 | 4.0333 | 1.4833 |  | 4.6667 | 4.8333 | 4 | 4.5667 | 4.6667 | 4.3333 | 4.4 | 4.3333 | 3.8314 | 2.8333 | 4.5333 | 4.3167 | 4.0725 | 5 | 8.1167
| Median |  | 4 |  | 119 | 4.05 | 1.4 |  | 5 | 5 | 4 | 4.55 | 4.7 | 4.4 | 4.4 | 4.4 | 3.7184 | 3 | 4.55 | 4.3 | 4.057 | 5 | 8.1
| Std Dev. |  | 0 |  |  | 0.12111 | 0.44008 |  | 0.5164 | 1.169 | 0 | 0.13663 | 0.23381 | 0.25033 | 0.26077 | 0.19664 | 0.32734 | 0.40825 | 0.5785 | 0.098319 | 0.1269 | 0 | 0.26394
| Rel.Std.Dev. |  | 0% |  |  | 3.003% | 29.67% |  | 11.07% | 24.19% | 0% | 2.992% | 5.01% | 5.777% | 5.927% | 4.538% | 8.543% | 14.41% | 12.76% | 2.278% | 3.116% | 0% | 3.252%
| PDM3 |  | -6.79% |  | 2673% | -6.013% | -65.43% |  | 8.745% | 12.63% | -6.79% | 6.415% | 8.745% | 0.9773% | 2.531% | 0.9773% | -10.72% | -33.98% | 5.638% | 0.5889% | -5.101% | 16.51% | 89.14%
| Z-Score (Absolute) |  | NA | 1.35 | 50.68 | 0.71 | 6.14 | 1.35 | NA | NA | NA | 0.43 | 0.64 | 0.07 | 0.07 | 0.07 | 1.14 | 3.27 | 0.35 | 0.11 | 0.63 | NA | 7.99
Results from laboratories 1, 8, 9, 11 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Y2O3 in OREAS 471 (Certified Value 31.7 ppm) | [28 empty] | Certified
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 1
2
| 1 | 1 | 30.4776 | 35.1332 | 26 | 32.0015 | 32.5729 | 33.0174 | 30.2236 | 31.8745 | 30.6046 | 32.1285 | 33.3984 | 30.8586 | 32.1285 | 35.68 | 30.4776 | 30.0507 | 32.0015 | 33.9063 | 30.3506 | 31.6053 | 36.9829 | 23.1122 | 34.3009 | 32.17 | 34.4919 | 32.1285 |  | 1
| 1 | 2 | 30.4776 | 34.1754 | 27 | 30.7316 | 32.1285 | 34.2873 | 29.5887 | 32.1285 | 31.2395 | 31.3665 | 34.5413 | 30.8586 | 31.7475 | 32.76 | 32.5094 | 30.4537 | 31.4935 | 33.1444 | 29.5887 | 32.1526 | 37.1998 | 21.0803 | 34.626 | 32.85 | 33.6551 | 31.7475 |  | #N/A
| 1 | 3 | 29.2077 | 36.2769 | 25 | 31.2395 | 32.2047 | 31.7475 | 28.8267 | 31.4935 | 34.9222 | 31.3665 | 33.1444 | 31.7475 | 29.8426 | 31.87 | 30.6046 | 30.0886 | 32.5094 | 32.8904 | 30.7316 | 32.1983 | 34.6575 | 29.3347 | 34.7722 | 32.57 | 35.1995 | 31.8745 |  | 16
| 1 | 4 | 29.2077 | 35.6101 | 24 | 31.8745 | 30.8078 | 33.0174 | 28.6997 | 32.3824 | 28.3188 | 32.0015 | 33.3984 | 32.5094 | 32.2555 | 32.26 | 30.7316 | 29.6904 | 32.8904 | 33.1444 | 30.4776 | 31.6713 | 37.6567 | 31.8745 | 34.0751 | 32.49 | 35.1136 | 31.6205 |  | 31.737
| 1 | 5 | 30.4776 | 36.3676 | 24 | 31.1126 | 32.3824 | 31.7475 | 28.9537 | 32.7634 | 31.6205 | 32.5094 | 33.6523 | 30.7316 | 30.6046 | 30.86 | 31.6205 | 30.7624 | 32.0015 | 34.0333 | 30.9856 | 31.4643 | 35.7625 | 32.0015 | 34.2009 | 32.31 | 34.3337 | 31.4935 |  | 59
| 1 | 6 | 27.9378 | 34.1365 | 26 | 31.8745 | 31.8745 | 31.7475 | 28.5727 | 32.5094 | 30.6046 | 30.7316 | 32.5094 | 31.8745 | 30.9856 | 31.62 | 33.6523 | 30.6161 | 31.6205 | 33.7793 | 30.2236 | 31.4072 | 39.4808 | 25.906 | 34.4514 | 32.06 | 35.024 | 32.3824
| Mean |  | 29.631 | 35.2833 | 25.3333 | 31.4724 | 31.9951 | 32.5941 | 29.1442 | 32.192 | 31.2184 | 31.684 | 33.4407 | 31.43 | 31.2607 | 32.5083 | 31.5993 | 30.277 | 32.0861 | 33.483 | 30.3929 | 31.7498 | 36.9567 | 27.2182 | 34.4044 | 32.4083 | 34.6363 | 31.8745
| Median |  | 29.8426 | 35.3717 | 25.5 | 31.557 | 32.1666 | 32.3825 | 28.8902 | 32.2555 | 30.9221 | 31.684 | 33.3984 | 31.303 | 31.3665 | 32.065 | 31.176 | 30.2712 | 32.0015 | 33.4619 | 30.4141 | 31.6383 | 37.0914 | 27.6203 | 34.3761 | 32.4 | 34.758 | 31.811
| Std Dev. |  | 1.0369 | 0.98336 | 1.2111 | 0.51688 | 0.62768 | 1.0369 | 0.63622 | 0.45963 | 2.1467 | 0.64628 | 0.66554 | 0.72172 | 0.94832 | 1.6792 | 1.2684 | 0.40317 | 0.53073 | 0.47966 | 0.47966 | 0.34333 | 1.6513 | 4.589 | 0.2635 | 0.28833 | 0.59507 | 0.33115
| Rel.Std.Dev. |  | 3.499% | 2.787% | 4.781% | 1.642% | 1.962% | 3.181% | 2.183% | 1.428% | 6.876% | 2.04% | 1.99% | 2.296% | 3.034% | 5.165% | 4.014% | 1.332% | 1.654% | 1.433% | 1.578% | 1.081% | 4.468% | 16.86% | 0.7659% | 0.8897% | 1.718% | 1.039%
| PDM3 |  | -6.636% | 11.17% | -20.18% | -0.8338% | 0.8134% | 2.701% | -8.17% | 1.434% | -1.634% | -0.1669% | 5.368% | -0.9672% | -1.501% | 2.43% | -0.4337% | -4.6% | 1.1% | 5.502% | -4.235% | 0.04049% | 16.45% | -14.24% | 8.405% | 2.115% | 9.135% | 0.4333%
| Z-Score (Absolute) |  | 1.99 | 3.16 | 5.91 | 0.31 | 0.17 | 0.71 | 2.43 | 0.35 | 0.54 | 0.12 | 1.48 | 0.35 | 0.5 | 0.64 | 0.19 | 1.4 | 0.25 | 1.52 | 1.29 | 0.06 | 4.69 | 4.19 | 2.36 | 0.54 | 2.57 | 0.06
| Analytical results for Yb2O3 in OREAS 471 (Certified Value 2.46 ppm) | [28 empty] | Certified
Yb2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 2.5051 | 2.5329 | 2.3 | 2.619 | 2.3913 | 2.8581 | 2.5051 | 2.5051 | 2.3913 | 2.5051 | 2.3343 | 2.8468 | 2.7557 | 2.7329 | 2.4596 | 2.4654 | 2.619 | 1.9813 | 2.4596 | 2.4824 | 2.1524 | 2.619 | 2.2082 | 2.9378 | 2.2644 | 2.3343 |  | 1
| 1 | 2 | 2.3913 | 2.4744 | 2.3 | 2.619 | 2.5051 | 2.6646 | 2.5051 | 2.5051 | 2.0497 | 2.3913 | 2.414 | 2.5051 | 2.4368 | 2.5735 | 2.4027 | 2.559 | 2.1635 | 2.5393 | 2.5165 | 2.4254 | 2.0427 | 2.5051 | 2.218 | 2.7215 | 2.2608 | 2.3116 |  | #N/A
| 1 | 3 | 2.5051 | 2.4165 | 2.3 | 2.619 | 2.3913 | 2.6873 | 2.5051 | 2.3913 | 2.0497 | 2.5051 | 2.4596 | 2.5165 | 2.1863 | 2.5735 | 2.4482 | 2.4817 | 2.5051 | 2.3799 | 2.5848 | 2.4972 | 2.2588 | 2.5051 | 2.2558 | 2.7443 | 2.2257 | 2.3116 |  | 16
| 1 | 4 | 2.5051 | 2.4042 | < 2.3 | 2.619 | 2.3913 | 2.7329 | 2.619 | 2.3913 | 1.8219 | 2.619 | 2.3116 | 2.5621 | 2.3343 | 2.471 | 2.8468 | 2.4803 | 2.3913 | 2.3343 | 2.4482 | 2.5063 | 2.2738 | 2.5051 | 2.182 | 2.5848 | 2.2706 | 2.2774 |  | 2.4554
| 1 | 5 | 2.619 | 2.7072 | < 2.3 | 2.619 | 2.3913 | 2.6646 | 2.5051 | 2.7329 | 2.2774 | 2.5051 | 2.2205 | 2.5735 | 2.414 | 2.619 | 2.4824 | 2.4879 | 2.5051 | 2.4596 | 2.5621 | 2.5188 | 2.2721 | 2.5051 | 2.2073 | 2.7101 | 2.2809 | 2.266 |  | 60
| 1 | 6 | 2.619 | 2.5077 | 2.3 | 2.5051 | 2.3913 | 2.8126 | 2.5051 | 2.619 | 2.0497 | 2.619 | 2.4254 | 2.5848 | 2.414 | 2.3571 | 2.471 | 2.5124 | 2.2774 | 2.6076 | 2.5279 | 2.4767 | 2.2086 | 2.619 | 2.2056 | 2.6873 | 2.1589 | 2.3002
| Mean |  | 2.5241 | 2.5071 | 2.3 | 2.6 | 2.4102 | 2.7367 | 2.5241 | 2.5241 | 2.1066 | 2.5241 | 2.3609 | 2.5981 | 2.4235 | 2.5545 | 2.5184 | 2.4978 | 2.4102 | 2.3837 | 2.5165 | 2.4845 | 2.2014 | 2.5431 | 2.2128 | 2.731 | 2.2436 | 2.3002
| Median |  | 2.5051 | 2.4911 | 2.3 | 2.619 | 2.3913 | 2.7101 | 2.5051 | 2.5051 | 2.0497 | 2.5051 | 2.3742 | 2.5678 | 2.414 | 2.5735 | 2.4653 | 2.4848 | 2.4482 | 2.4197 | 2.5222 | 2.4898 | 2.2337 | 2.5051 | 2.2077 | 2.7158 | 2.2626 | 2.3059
| Std Dev. |  | 0.085718 | 0.11004 | 0 | 0.046487 | 0.046487 | 0.081691 | 0.046487 | 0.13312 | 0.20049 | 0.085718 | 0.088984 | 0.12586 | 0.18703 | 0.12856 | 0.16318 | 0.03368 | 0.16761 | 0.22112 | 0.054372 | 0.032761 | 0.090735 | 0.058802 | 0.024219 | 0.11555 | 0.045492 | 0.024948
| Rel.Std.Dev. |  | 3.396% | 4.389% | 0% | 1.788% | 1.929% | 2.985% | 1.842% | 5.274% | 9.517% | 3.396% | 3.769% | 4.844% | 7.717% | 5.033% | 6.48% | 1.348% | 6.954% | 9.276% | 2.161% | 1.319% | 4.122% | 2.312% | 1.095% | 4.231% | 2.028% | 1.085%
| PDM3 |  | 2.799% | 2.107% | -6.329% | 5.89% | -1.839% | 11.46% | 2.799% | 2.799% | -14.21% | 2.799% | -3.848% | 5.813% | -1.298% | 4.035% | 2.567% | 1.727% | -1.839% | -2.921% | 2.49% | 1.183% | -10.34% | 3.572% | -9.88% | 11.22% | -8.628% | -6.322%
| Z-Score (Absolute) |  | 0.16 | 0.03 | 4.42 | 0.71 | 0.67 | 1.71 | 0.16 | 0.16 | 2.89 | 0.16 | 1.04 | 0.7 | 0.58 | 0.38 | 0.12 | 0.03 | 0.67 | 0.87 | 0.1 | 0.13 | 2.2 | 0.3 | 2.12 | 1.67 | 1.89 | 1.48
| Analytical results for Zn in OREAS 471 (Certified Value 78 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 10 | 17 | 18 | 21 | 23 | 24 | 26 | 27 | [12 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | BF*MS | PF*OES | BF*MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | [12 empty] | 0
1
| 1 | 1 | 80 | 85.0209 | 100 | 80 | 80 | 85 | 80 | 80 | 74.5331 | < 100 | 74.426 | 71 | 79.52 | 76.7025 | 77 | [12 empty] | 1
| 1 | 2 | 80 | 81.9899 | 126 | 75 | 83 | 90 | 70 | 70 | 75.4227 | < 100 | 84.798 | 69 | 78.43 | 80.09 | 77 | [12 empty] | #N/A
| 1 | 3 | 80 | 86.4973 | 80 | 78 | 79 | 88 | 80 | 75 | 74.6375 | < 100 | 76.568 | 71 | 78.18 | 75.132 | 76 | [12 empty] | 16
| 1 | 4 | 70 | 90.6653 | 119 | 79 | 83 | 83 | 80 | 60 | 72.6089 | < 100 | 82.571 | 71 | 79.33 | 80.811 | 77 | [12 empty] | 78.3844
| 1 | 5 | 70 | 84.4965 | 81 | 80 | 82 | 87 | 80 | 75 | 77.6823 | < 100 | 75.507 | 69 | 79.04 | 77.07 | 77 | [12 empty] | 61
| 1 | 6 | 70 | 84.1951 | 96 | 77 | 91 | 83 | 80 | 65 | 77.9175 | < 100 | 74.03 | 74 | 79.32 | 83.0025 | 79
| Mean |  | 75 | 85.4775 | 100.3333 | 78.1667 | 83 | 86 | 78.3333 | 70.8333 | 75.467 |  | 77.9833 | 70.8333 | 78.97 | 78.8013 | 77.1667
| Median |  | 75 | 84.7587 | 98 | 78.5 | 82.5 | 86 | 80 | 72.5 | 75.0301 |  | 76.0375 | 71 | 79.18 | 78.58 | 77
| Std Dev. |  | 5.4772 | 2.9303 | 19.0438 | 1.9408 | 4.2426 | 2.8284 | 4.0825 | 7.3598 | 2.0319 |  | 4.5586 | 1.8348 | 0.54318 | 2.9737 | 0.98319
| Rel.Std.Dev. |  | 7.303% | 3.428% | 18.98% | 2.483% | 5.112% | 3.289% | 5.212% | 10.39% | 2.692% |  | 5.846% | 2.59% | 0.6878% | 3.774% | 1.274%
| PDM3 |  | -4.318% | 9.049% | 28% | -0.2778% | 5.888% | 9.716% | -0.06513% | -9.633% | -3.722% |  | -0.5117% | -9.633% | 0.7471% | 0.5319% | -1.554%
| Z-Score (Absolute) |  | NA | 1.02 | 3.09 | 0 | 0.67 | 1.09 | NA | 1.02 | 0.38 | 3.93 | 0.03 | 1.02 | 0.11 | 0.09 | 0.14
Results from laboratories 1 and 8 were removed due to their 10 ppm reading resolution.
| Analytical results for Zr in OREAS 471 (Certified Value 316 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 27 | [7 empty] | ppm
|  |  | BF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | BF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 328 | 318.3455 | 263 | 280 | 311.6 | 352 | 299 | 317 | 329 | 326 | 312 | 355 | 331 | 303.1597 | 326 | 315.7 | 325.4 | 151.934 | 328 | 240.2 | [7 empty] | 1
| 1 | 2 | 325 | 319.55 | 288 | 263 | 317.9 | 341 | 303 | 319 | 333 | 351 | 300 | 335 | 336 | 316.3257 | 320 | 324.6 | 320.3 | 159.863 | 329 | 238.1 | [7 empty] | #N/A
| 1 | 3 | 319 | 313.2725 | 261 | 283 | 307.3 | 358 | 315 | 320 | 326 | 330 | 292 | 324 | 324 | 305.2552 | 320 | 307.7 | 333.5 | 162.44 | 277 | 237.3 | [7 empty] | 16
| 1 | 4 | 320 | 310.3285 | 235 | 272 | 302.8 | 330 | 311 | 317 | 331 | 325 | 305 | 344 | 341 | 298.7697 | 323 | 314.2 | 327.7 | 155.819 | 304 | 243.3 | [7 empty] | 315.5706
| 1 | 5 | 321 | 312.5864 | 273 | 274 | 302.8 | 351 | 310 | 320 | 345 | 330 | 304 | 330 | 328 | 313.0913 | 324 | 313.8 | 334.1 | 157.991 | 308 | 239 | [7 empty] | 62
| 1 | 6 | 330 | 311.3828 | 279 | 265 | 315.5 | 343 | 309 | 317 | 319 | 326 | 297 | 319 | 335 | 312.362 | 329 | 318 | 325.3 | 149.573 | 242 | 242.9
| Mean |  | 323.8333 | 314.2443 | 266.5 | 272.8333 | 309.65 | 345.8333 | 307.8333 | 318.3333 | 330.5 | 331.3333 | 301.6667 | 334.5 | 332.5 | 308.1606 | 323.6667 | 315.6667 | 327.7167 | 156.27 | 298 | 240.1333
| Median |  | 323 | 312.9294 | 268 | 273 | 309.45 | 347 | 309.5 | 318 | 330 | 328 | 302 | 332.5 | 333 | 308.8086 | 323.5 | 314.95 | 326.55 | 156.905 | 306 | 239.6
| Std Dev. |  | 4.535 | 3.7996 | 18.4147 | 7.9352 | 6.4102 | 9.9482 | 5.8109 | 1.5055 | 8.6197 | 9.8725 | 6.9474 | 13.2778 | 6.091 | 6.7863 | 3.5024 | 5.5561 | 5.2992 | 4.8538 | 33.3886 | 2.4953
| Rel.Std.Dev. |  | 1.4% | 1.209% | 6.91% | 2.908% | 2.07% | 2.877% | 1.888% | 0.4729% | 2.608% | 2.98% | 2.303% | 3.969% | 1.832% | 2.202% | 1.082% | 1.76% | 1.617% | 3.106% | 11.2% | 1.039%
| PDM3 |  | 2.618% | -0.4203% | -15.55% | -13.54% | -1.876% | 9.59% | -2.452% | 0.8755% | 4.731% | 4.995% | -4.406% | 5.998% | 5.365% | -2.348% | 2.566% | 0.03043% | 3.849% | -50.48% | -5.568% | -23.91%
| Z-Score (Absolute) |  | 0.42 | 0.03 | 2.27 | 1.97 | 0.25 | 1.44 | 0.33 | 0.16 | 0.73 | 0.77 | 0.62 | 0.91 | 0.82 | 0.32 | 0.41 | 0.03 | 0.6 | 7.42 | 0.79 | 3.5
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 471 (Certified Value 16.41 wt.%) | [14 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 16.11 | 16.407 | 15.9 | 16.18 | 16.4883 | 16.7 | 16.6214 | 16.69 | 16.8549 | 15.47 | 16.56 | 16.45 | 15.626 | 1
| 1 | 2 | 16.38 | 16.339 | 16.01 | 16.18 | 16.6145 | 16.2 | 16.586 | 16.57 | 16.7804 | 15.3101 | 16.55 | 16.47 | 15.985 | #N/A
| 1 | 3 | 16.18 | 16.377 | 16.1 | 16.18 | 16.4943 | 16.6 | 16.596 | 16.54 | 16.7824 | 15.4518 | 16.54 | 16.46 | 15.928 | 16
| 1 | 4 | 16.25 | 16.373 | 16.12 | 16.15 | 16.4256 | 16.7 | 16.6162 | 16.56 | 16.821 | 15.3761 | 16.54 | 16.43 | 15.834 | 16.4088
| 1 | 5 | 16.31 | 16.4 | 16.24 | 16.07 | 16.5266 | 16.6 | 16.6043 | 16.61 | 16.7971 | 15.5236 | 16.56 | 16.44 | 16.042 | 64
| 1 | 6 | 16.3 | 16.404 | 16.22 | 16.17 | 16.4261 | 16.6 | 16.5927 | 16.56 | 16.7852 | 15.4859 | 16.55 | 16.48 | 15.966
| Mean |  | 16.255 | 16.3833 | 16.0983 | 16.155 | 16.4959 | 16.5667 | 16.6028 | 16.5883 | 16.8035 | 15.4363 | 16.55 | 16.455 | 15.8968
| Median |  | 16.275 | 16.3885 | 16.11 | 16.175 | 16.4913 | 16.6 | 16.6002 | 16.565 | 16.7911 | 15.4609 | 16.55 | 16.455 | 15.947
| Std Dev. |  | 0.097314 | 0.02599 | 0.12844 | 0.043243 | 0.070536 | 0.18619 | 0.013831 | 0.054924 | 0.029345 | 0.07875 | 0.0089443 | 0.018708 | 0.1496
| Rel.Std.Dev. |  | 0.5987% | 0.1586% | 0.7978% | 0.2677% | 0.4276% | 1.124% | 0.08331% | 0.3311% | 0.1746% | 0.5102% | 0.05404% | 0.1137% | 0.9411%
| PDM3 |  | -0.937% | -0.1549% | -1.892% | -1.546% | 0.5312% | 0.9624% | 1.183% | 1.094% | 2.406% | -5.927% | 0.8608% | 0.2818% | -3.12%
| Z-Score (Absolute) |  | 0.91 | 0.33 | 1.63 | 1.37 | 0.19 | 0.51 | 0.67 | 0.61 | 1.59 | 4.65 | 0.43 | 0 | 2.55
| Analytical results for As2O3 in OREAS 471 (Indicative Value 18.5 ppm) | [14 empty] | Indicative
As2O3 | Round | Replicate | Lab | Lab | Lab | [10 empty] | 1
| No | No | 05 | 23 | 27 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 1
1
| 1 | 1 | 10 | < 100 | 11.9 | [10 empty] | 1
| 1 | 2 |  | < 100 | 13.2 | [10 empty] | 5
| 1 | 3 |  | < 100 | 14.5 | [10 empty] | 16
| 1 | 4 |  | < 100 | 14.5 | [10 empty] | 18.4882
| 1 | 5 | 30 | < 100 | 14.5 | [10 empty] | 11
| 1 | 6 | 30 | < 100 | 13.2
| Mean |  | 23.3333 |  | 13.6333
| Median |  | 30 |  | 13.85
| Std Dev. |  | 11.547 |  | 1.0614
| Rel.Std.Dev. |  | 49.49% |  | 7.786%
| PDM3 |  | 26.21% |  | -26.26%
| Z-Score (Absolute) |  | 0 | 1.86 | 0.67
| Analytical results for Au in OREAS 471 (Indicative Value < 20 ppm) | [14 empty] | Indicative
Au | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 2
2
| 1 | 1 | < 20 | [12 empty] | 1
| 1 | 2 | < 20 | [12 empty] | 6
| 1 | 3 | < 20 | [12 empty] | 16
| 1 | 4 | < 20 | [12 empty] | < 20
| 1 | 5 | < 20 | [12 empty] | 12
| 1 | 6 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 471 (Certified Value 0.272 wt.%) | [14 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 23 | 24 | 27 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.294 | 0.24 | 0.26 | 0.2625 | 0.3 | 0.27308 | 0.268 | 0.28 | 0.28 | 0.123 |  |  |  | 1
| 1 | 2 | 0.2835 | 0.25 | 0.26 | 0.262 | 0.28 | 0.27667 | 0.268 | 0.28 | 0.29 | 0.123 |  |  |  | #N/A
| 1 | 3 | 0.2835 | 0.24 | 0.26 | 0.2565 | 0.3 | 0.27538 | 0.257 | 0.28 | 0.28 | 0.123 |  |  |  | 16
| 1 | 4 | 0.29295 | 0.25 | 0.26 | 0.2585 | 0.32 | 0.2742 | 0.257 | 0.28 | 0.27 | 0.123 |  |  |  | 0.2722
| 1 | 5 | 0.2856 | 0.25 | 0.25 | 0.265 | 0.3 | 0.27375 | 0.268 | 0.28 | 0.27 | 0.123 |  |  |  | 65
| 1 | 6 | 0.28875 | 0.24 | 0.25 | 0.263 | 0.3 | 0.27425 | 0.268 | 0.28 | 0.29 | 0.134
| Mean |  | 0.28805 | 0.245 | 0.25667 | 0.26125 | 0.3 | 0.27455 | 0.26433 | 0.28 | 0.28 | 0.12483
| Median |  | 0.28718 | 0.245 | 0.26 | 0.26225 | 0.3 | 0.27422 | 0.268 | 0.28 | 0.28 | 0.123
| Std Dev. |  | 0.0046327 | 0.0054772 | 0.005164 | 0.0031425 | 0.012649 | 0.0012815 | 0.0056804 | 0 | 0.0089443 | 0.0044907
| Rel.Std.Dev. |  | 1.608% | 2.236% | 2.012% | 1.203% | 4.216% | 0.4667% | 2.149% | 0% | 3.194% | 3.597%
| PDM3 |  | 5.825% | -9.991% | -5.705% | -4.021% | 10.21% | 0.8662% | -2.888% | 2.867% | 2.867% | -54.14%
| Z-Score (Absolute) |  | 1.08 | 1.41 | 0.74 | 0.47 | 1.77 | 0.3 | 0.3 | 0.61 | 0.61 | 8.36
| Analytical results for Bi2O3 in OREAS 471 (Indicative Value < 11 ppm) | [14 empty] | Indicative
Bi2O3 | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 100 | < 11 | [11 empty] | 1
| 1 | 2 | < 100 | < 11 | [11 empty] | 7
| 1 | 3 | < 100 | < 11 | [11 empty] | 16
| 1 | 4 | < 100 | < 11 | [11 empty] | < 11
| 1 | 5 | < 100 | < 11 | [11 empty] | 13
| 1 | 6 | < 100 | < 11
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for CaO in OREAS 471 (Certified Value 1.21 wt.%) | [14 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.18 | 1.2193 | 1.21 | 1.2 | 1.1814 | 1.23 | 1.1987 | 1.24 | 1.2901 | 1.1685 | 1.22 | 1.2 | 1.511 | 1
| 1 | 2 | 1.2 | 1.219 | 1.23 | 1.2 | 1.1925 | 1.22 | 1.1996 | 1.23 | 1.2884 | 1.1614 | 1.21 | 1.19 | 1.651 | #N/A
| 1 | 3 | 1.2 | 1.2089 | 1.23 | 1.19 | 1.1867 | 1.21 | 1.2027 | 1.22 | 1.2938 | 1.1788 | 1.22 | 1.2 | 1.679 | 16
| 1 | 4 | 1.2 | 1.217 | 1.24 | 1.19 | 1.1798 | 1.21 | 1.204 | 1.22 | 1.2911 | 1.1767 | 1.22 | 1.21 | 1.707 | 1.2056
| 1 | 5 | 1.2 | 1.2129 | 1.24 | 1.19 | 1.1861 | 1.19 | 1.2018 | 1.24 | 1.3036 | 1.1948 | 1.21 | 1.21 | 1.665 | 66
| 1 | 6 | 1.21 | 1.216 | 1.25 | 1.2 | 1.1851 | 1.21 | 1.1998 | 1.23 | 1.3006 | 1.1766 | 1.21 | 1.19 | 1.651
| Mean |  | 1.1983 | 1.2155 | 1.2333 | 1.195 | 1.1853 | 1.2117 | 1.2011 | 1.23 | 1.2946 | 1.1761 | 1.215 | 1.2 | 1.644
| Median |  | 1.2 | 1.2165 | 1.235 | 1.195 | 1.1856 | 1.21 | 1.2008 | 1.23 | 1.2925 | 1.1766 | 1.215 | 1.2 | 1.658
| Std Dev. |  | 0.0098319 | 0.0039997 | 0.013663 | 0.0054772 | 0.0044805 | 0.013292 | 0.0020371 | 0.0089443 | 0.0061513 | 0.011238 | 0.0054772 | 0.0089443 | 0.068443
| Rel.Std.Dev. |  | 0.8205% | 0.3291% | 1.108% | 0.4583% | 0.378% | 1.097% | 0.1696% | 0.7272% | 0.4752% | 0.9555% | 0.4508% | 0.7454% | 4.163%
| PDM3 |  | -0.6008% | 0.8242% | 2.302% | -0.8773% | -1.685% | 0.5052% | -0.3711% | 2.026% | 7.384% | -2.442% | 0.7817% | -0.4625% | 36.37%
| Z-Score (Absolute) |  | 0.54 | 0.16 | 0.88 | 0.67 | 1.07 | 0 | 0.43 | 0.74 | 3.36 | 1.44 | 0.13 | 0.47 | 17.49
| Analytical results for Cd in OREAS 471 (Indicative Value < 10 ppm) | [14 empty] | Indicative
Cd | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 10 | [12 empty] | 1
| 1 | 2 | < 10 | [12 empty] | 8
| 1 | 3 | < 10 | [12 empty] | 16
| 1 | 4 | < 10 | [12 empty] | < 10
| 1 | 5 | < 10 | [12 empty] | 14
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for CeO2 in OREAS 471 (Certified Value 243 ppm) | [14 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 03 | 05 | 06 | 12 | 13 | 15 | 16 | 17 | 22 | 23 | 24 | 27 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 100 | 236 | 400 | 300 | 200 | 300 | 300 | 237.1056 | 193.2697 | 300 | 303.5372 | 206.4 |  | 1
| 1 | 2 | 200 | 224.2 | 300 | 300 | 200 | 300 | 300 | 226.1685 | 191.1292 | 200 | 302.7084 | 203.9 |  | #N/A
| 1 | 3 | 300 | 283.2 | 100 | 300 | 100 | 300 | 300 | 204.0798 | 193.5593 | 100 | 305.085 | 197.8 |  | 16
| 1 | 4 | 400 | 253.7 | 200 | 300 | 200 | 300 | 300 | 206.5942 | 208.8282 | 200 | 305.3845 | 202.7 |  | 243.3942
| 1 | 5 | 200 | 218.3 | 300 | 300 | < 100 | 300 | 200 | 224.5385 | 196.612 | 200 | 304.5058 | 210.1 |  | 67
| 1 | 6 | 300 | 230.1 | 100 | 200 | 200 | 300 | 300 | 199.7468 | 189.1291 | 200 | 304.6456 | 201.5
| Mean |  | 280 | 240.9167 | 233.3333 | 283.3333 | 180 | 300 | 283.3333 | 216.3722 | 195.4213 | 200 | 304.3111 | 203.7333
| Median |  | 300 | 233.05 | 250 | 300 | 200 | 300 | 300 | 215.5663 | 193.4145 | 200 | 304.5757 | 203.3
| Std Dev. |  | 83.666 | 24.0143 | 121.106 | 40.8248 | 44.7214 | 0 | 40.8248 | 14.9375 | 7.0324 | 63.2456 | 1.0068 | 4.2174
| Rel.Std.Dev. |  | 29.88% | 9.968% | 51.9% | 14.41% | 24.85% | 0% | 14.41% | 6.904% | 3.599% | 31.62% | 0.3308% | 2.07%
| PDM3 |  | 15.04% | -1.018% | -4.134% | 16.41% | -26.05% | 23.26% | 16.41% | -11.1% | -19.71% | -17.83% | 25.03% | -16.29%
| Z-Score (Absolute) |  | 0.08 | 0.06 | 0.06 | 0.79 | 1.35 | 1.08 | 0.79 | 0.36 | 0.71 | 0.64 | 1.15 | 0.57
| Analytical results for Cl in OREAS 471 (Indicative Value 0.238 wt.%) | [14 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | wt.%
|  |  | BF*XRF | BF*XRF | [11 empty] | 3
3
| 1 | 1 | 0.07 | 0.38 | [11 empty] | 1
| 1 | 2 | 0.07 | 0.4 | [11 empty] | 9
| 1 | 3 | 0.07 | 0.41 | [11 empty] | 16
| 1 | 4 | 0.07 | 0.42 | [11 empty] | 0.2375
| 1 | 5 | 0.07 | 0.41 | [11 empty] | 15
| 1 | 6 | 0.08 | 0.4
| Mean |  | 0.071667 | 0.40333
| Median |  | 0.07 | 0.405
| Std Dev. |  | 0.0040825 | 0.013663
| Rel.Std.Dev. |  | 5.696% | 3.387%
| PDM3 |  | -69.82% | 69.82%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co3O4 in OREAS 471 (Indicative Value < 50 ppm) | [14 empty] | Indicative
Co3O4 | Round | Replicate | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 23 | 27 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 50 | < 100 | < 14 | [10 empty] | 1
| 1 | 2 | < 50 | < 100 | < 14 | [10 empty] | 10
| 1 | 3 | < 50 | < 100 | < 14 | [10 empty] | 16
| 1 | 4 | < 50 | < 100 | < 14 | [10 empty] | < 50
| 1 | 5 | < 50 | < 100 | < 14 | [10 empty] | 16
| 1 | 6 | < 50 | < 100 | < 14
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Cr2O3 in OREAS 471 (Indicative Value 78 ppm) | [14 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 17 | 20 | 23 | 24 | 27 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 100 | 80 | < 50 | 200 | 80 | 74.8029 | < 100 | < 100 | < 100 | < 150 |  |  |  | 1
| 1 | 2 | 100 | 80 | 60 | 200 | 80 | 72.908 | < 100 | < 100 | < 100 | < 150 |  |  |  | 11
| 1 | 3 | 100 | 100 | 60 | 300 | 95 | 75.7309 | < 100 | < 100 | < 100 | < 150 |  |  |  | 16
| 1 | 4 | 100 | 90 | 60 | < 100 | 85 | 79.5319 | < 100 | < 100 | < 100 | < 150 |  |  |  | 78.1672
| 1 | 5 | 100 | 90 | < 50 | 100 | 90 | 68.1767 | < 100 | < 100 | < 100 | < 150 |  |  |  | 17
| 1 | 6 | 100 | 90 | < 50 | 200 | 85 | 73.8651 | < 100 | < 100 | < 100 | < 150
| Mean |  | 100 | 88.3333 | 60 | 200 | 85.8333 | 74.1692
| Median |  | 100 | 90 | 60 | 200 | 85 | 74.334
| Std Dev. |  | 0 | 7.5277 | 0 | 70.7107 | 5.8452 | 3.72
| Rel.Std.Dev. |  | 0% | 8.522% | 0% | 35.36% | 6.81% | 5.016%
| PDM3 |  | 27.93% | 13.01% | -23.24% | 155.9% | 9.807% | -5.115%
| Z-Score (Absolute) |  | 0.75 | 0.42 | 0.89 | 2.89 | 0.35 | 0.02 | 0.67 | 0.67 | 0.67 | 0.02
| Analytical results for Cs2O in OREAS 471 (Indicative Value < 11 ppm) | [14 empty] | Indicative
Cs2O | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 100 | < 11 | [11 empty] | 1
| 1 | 2 | < 100 | < 11 | [11 empty] | 12
| 1 | 3 | < 100 | < 11 | [11 empty] | 16
| 1 | 4 | < 100 | < 11 | [11 empty] | < 11
| 1 | 5 | < 100 | < 11 | [11 empty] | 18
| 1 | 6 | < 100 | < 11
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for CuO in OREAS 471 (Indicative Value 52 ppm) | [14 empty] | Indicative
CuO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 17 | 23 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 50 | < 200 | < 100 | 60 | 38.3324 | < 100 | 65.1 | [6 empty] | 1
| 1 | 2 | < 50 | < 200 | < 100 | 55 | 31.1135 | < 100 | 75.1 | [6 empty] | 13
| 1 | 3 | < 50 | < 200 | < 100 | 70 | 33.1681 | < 100 | 71.4 | [6 empty] | 16
| 1 | 4 | 50 | < 200 | < 100 | 70 | 35.2601 | < 100 | 78.9 | [6 empty] | 52.2216
| 1 | 5 | 50 | < 200 | < 100 | 60 | 64.4463 | < 100 | 78.9 | [6 empty] | 19
| 1 | 6 | 50 | < 200 | < 100 | 65 | 35.1957 | < 100 | 71.4
| Mean |  | 50 |  |  | 63.3333 | 39.586 |  | 73.4667
| Median |  | 50 |  |  | 62.5 | 35.2279 |  | 73.25
| Std Dev. |  | 0 |  |  | 6.0553 | 12.4142 |  | 5.2963
| Rel.Std.Dev. |  | 0% |  |  | 9.561% | 31.36% |  | 7.209%
| PDM3 |  | -4.254% |  |  | 21.28% | -24.2% |  | 40.68%
| Z-Score (Absolute) |  | 0.67 | 2.7 | 0 | 0.72 | 0.56 | 0 | 1.26
| Analytical results for Dy2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Dy2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 22 | 23 | 24 | 27 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 5.3285 | < 23 |  |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 5.2358 | < 23 |  |  |  |  | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 5.256 | < 23 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 5.2421 | < 23 |  |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | < 100 | 5.2754 | < 23 |  |  |  |  | 68
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 5.1883 | < 23
| Mean |  |  |  | 100 |  |  |  |  | 5.2544
| Median |  |  |  | 100 |  |  |  |  | 5.249
| Std Dev. | [8 empty] | 0.046483
| Rel.Std.Dev. | [8 empty] | 0.8847%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Er2O3 in OREAS 471 (Indicative Value 105 ppm) | [14 empty] | Indicative
Er2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 24 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | 100 | 2.7252 | 178.4 | [6 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | 100 | 2.8084 | 170.4 | [6 empty] | 14
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | 200 | 2.8279 | 179.5 | [6 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | 200 | 2.72 | 183 | [6 empty] | 104.5107
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | 100 | 2.7691 | 180.7 | [6 empty] | 20
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | 100 | 2.7435 | 172.7
| Mean | [5 empty] | 133.3333 | 2.7657 | 177.45
| Median | [5 empty] | 100 | 2.7563 | 178.95
| Std Dev. | [5 empty] | 51.6398 | 0.044536 | 4.8731
| Rel.Std.Dev. | [5 empty] | 38.73% | 1.61% | 2.746%
| PDM3 | [5 empty] | 27.58% | -97.35% | 69.79%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Eu2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Eu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 23 | 24 | 27 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 3.9494 | 383.3 | [5 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 4.0328 | 402.9 | [5 empty] | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 3.9579 | 399.5 | [5 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 3.9475 | 406.4 | [5 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 4.013 | 375.2 | [5 empty] | 69
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 3.9637 | 390.2
| Mean | [7 empty] | 3.9774 | 392.9167
| Median | [7 empty] | 3.9608 | 394.85
| Std Dev. | [7 empty] | 0.036281 | 12.1536
| Rel.Std.Dev. | [7 empty] | 0.9122% | 3.093%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for F in OREAS 471 (Indicative Value < 5000 ppm) | [14 empty] | Indicative
F | Round | Replicate | Lab | [12 empty] | 1
| No | No | 23 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 5000 | [12 empty] | 1
| 1 | 2 | < 5000 | [12 empty] | 15
| 1 | 3 | < 5000 | [12 empty] | 16
| 1 | 4 | < 5000 | [12 empty] | < 5000
| 1 | 5 | < 5000 | [12 empty] | 21
| 1 | 6 | < 5000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 471 (Certified Value 6.2 wt.%) | [14 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 6.26 | 6.1358 | 6.19 | 6.12 | 6.0353 | 6.34 | 6.162 | 6.13 | 6.2315 | 6.2671 | 6.15 | 6.35 | 7.72 | 1
| 1 | 2 | 6.34 | 6.1493 | 6.27 | 6.14 | 6.0616 | 6.17 | 6.1253 | 6.07 | 6.2536 | 6.2429 | 6.17 | 6.33 | 8.535 | #N/A
| 1 | 3 | 6.29 | 6.1891 | 6.29 | 6.12 | 6.0461 | 6.26 | 6.143 | 6 | 6.2266 | 6.1201 | 6.14 | 6.34 | 8.664 | 16
| 1 | 4 | 6.28 | 6.1791 | 6.31 | 6.1 | 6.0301 | 6.33 | 6.1638 | 6.01 | 6.2732 | 6.2658 | 6.14 | 6.36 | 8.907 | 6.2031
| 1 | 5 | 6.27 | 6.19 | 6.35 | 6.05 | 6.044 | 6.33 | 6.1395 | 6.08 | 6.2714 | 6.2869 | 6.14 | 6.36 | 8.607 | 70
| 1 | 6 | 6.31 | 6.206 | 6.34 | 6.14 | 6.0487 | 6.36 | 6.1537 | 6.06 | 6.2626 | 6.4322 | 6.15 | 6.35 | 8.478
| Mean |  | 6.2917 | 6.1749 | 6.2917 | 6.1117 | 6.0443 | 6.2983 | 6.1479 | 6.0583 | 6.2531 | 6.2692 | 6.1483 | 6.3483 | 8.4852
| Median |  | 6.285 | 6.1841 | 6.3 | 6.12 | 6.0451 | 6.33 | 6.1484 | 6.065 | 6.2581 | 6.2664 | 6.145 | 6.35 | 8.571
| Std Dev. |  | 0.029269 | 0.026825 | 0.058109 | 0.033714 | 0.010972 | 0.071391 | 0.014761 | 0.047924 | 0.020001 | 0.099821 | 0.01169 | 0.01169 | 0.40318
| Rel.Std.Dev. |  | 0.4652% | 0.4344% | 0.9236% | 0.5516% | 0.1815% | 1.133% | 0.2401% | 0.791% | 0.3199% | 1.592% | 0.1901% | 0.1841% | 4.752%
| PDM3 |  | 1.427% | -0.4557% | 1.427% | -1.475% | -2.561% | 1.535% | -0.8908% | -2.334% | 0.8061% | 1.064% | -0.8836% | 2.341% | 36.79%
| Z-Score (Absolute) |  | 0.27 | 0.55 | 0.27 | 1 | 1.48 | 0.32 | 0.75 | 1.38 | 0 | 0.11 | 0.74 | 0.67 | 15.81
| Analytical results for Fr in OREAS 471 (Indicative Value < 30 ppm) | [14 empty] | Indicative
Fr | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 30 | [12 empty] | 1
| 1 | 2 | < 30 | [12 empty] | 16
| 1 | 3 | < 30 | [12 empty] | 16
| 1 | 4 | < 30 | [12 empty] | < 30
| 1 | 5 | < 30 | [12 empty] | 22
| 1 | 6 | < 30
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ga2O3 in OREAS 471 (Indicative Value 117 ppm) | [14 empty] | Indicative
Ga2O3 | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | 104.8 | [11 empty] | 1
| 1 | 2 | < 100 | 121 | [11 empty] | 17
| 1 | 3 | < 100 | 114.3 | [11 empty] | 16
| 1 | 4 | < 100 | 125 | [11 empty] | 116.7214
| 1 | 5 | < 100 | 116.9 | [11 empty] | 23
| 1 | 6 | < 100 | 118.3
| Mean |  |  | 116.7167
| Median |  |  | 117.6
| Std Dev. |  |  | 6.8898
| Rel.Std.Dev. |  |  | 5.903%
| PDM3 |  |  | -0.004027%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Gd2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 06 | 12 | 13 | 15 | 16 | 22 | 23 | 24 | 27 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 9.5079 | 134.9 |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 9.4779 | 146.4 |  |  |  | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 9.6137 | 153.3 |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 9.4965 | 152.1 |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 9.647 | 149.8 |  |  |  | 71
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 9.6302 | 151
| Mean | [9 empty] | 9.5622 | 147.9167
| Median | [9 empty] | 9.5608 | 150.4
| Std Dev. | [9 empty] | 0.075933 | 6.801
| Rel.Std.Dev. | [9 empty] | 0.7941% | 4.598%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Ge in OREAS 471 (Indicative Value < 30 ppm) | [14 empty] | Indicative
Ge | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 30 | [12 empty] | 1
| 1 | 2 | < 30 | [12 empty] | 18
| 1 | 3 | < 30 | [12 empty] | 16
| 1 | 4 | < 30 | [12 empty] | < 30
| 1 | 5 | < 30 | [12 empty] | 24
| 1 | 6 | < 30
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for H2O- in OREAS 471 (Indicative Value 1.1 wt.%) | [14 empty] | Indicative
H2O- | Round | Replicate | Lab | [12 empty] | 1
| No | No | 17 | [12 empty] | wt.%
|  |  | BF*XRF | [12 empty] | 2
2
| 1 | 1 | 1.1047 | [12 empty] | 1
| 1 | 2 | 1.1098 | [12 empty] | 19
| 1 | 3 | 1.1081 | [12 empty] | 16
| 1 | 4 | 1.1031 | [12 empty] | 1.1023
| 1 | 5 | 1.1034 | [12 empty] | 25
| 1 | 6 | 1.0849
| Mean |  | 1.1023
| Median |  | 1.104
| Std Dev. |  | 0.0089471
| Rel.Std.Dev. |  | 0.8117%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 471 (Indicative Value < 100 ppm) | [14 empty] | Indicative
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 21 | 23 | 27 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 200 | < 100 |  | < 50 | < 100 | < 12 | [7 empty] | 1
| 1 | 2 | < 200 | < 100 | 5 | < 50 | < 100 | < 12 | [7 empty] | 20
| 1 | 3 | < 200 | < 100 |  | < 50 | < 100 | < 12 | [7 empty] | 16
| 1 | 4 | < 200 | < 100 | 5 | < 50 | < 100 | < 12 | [7 empty] | < 100
| 1 | 5 | < 200 | < 100 | 10 | < 50 | < 100 | < 12 | [7 empty] | 26
| 1 | 6 | < 200 | < 100 |  | < 50 | < 100 | < 12
| Mean |  |  |  | 6.6667
| Median |  |  |  | 5
| Std Dev. |  |  |  | 2.8868
| Rel.Std.Dev. |  |  |  | 43.3%
PDM3
| Z-Score (Absolute) |  | 1.95 | 0.39 | 0.96 | 0.39 | 0.39 | 0.98
| Analytical results for Hg in OREAS 471 (Indicative Value < 20 ppm) | [14 empty] | Indicative
Hg | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 20 | [12 empty] | 1
| 1 | 2 | < 20 | [12 empty] | 5
| 1 | 3 | < 20 | [12 empty] | 16
| 1 | 4 | < 20 | [12 empty] | < 20
| 1 | 5 | < 20 | [12 empty] | 11
| 1 | 6 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Ho2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 24 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.9576 | 92.8 | [6 empty] | 1
| 1 | 2 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.94512 | 100.8 | [6 empty] | #N/A
| 1 | 3 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.91367 | 101.9 | [6 empty] | 16
| 1 | 4 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.92735 | 111.1 | [6 empty] | < 100
| 1 | 5 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.90248 | 106.5 | [6 empty] | 72
| 1 | 6 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.94163 | 100.8
| Mean | [6 empty] | 0.93131 | 102.3167
| Median | [6 empty] | 0.93449 | 101.35
| Std Dev. | [6 empty] | 0.020719 | 6.1642
| Rel.Std.Dev. | [6 empty] | 2.225% | 6.025%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for I in OREAS 471 (Indicative Value < 10 ppm) | [14 empty] | Indicative
I | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 10 | [12 empty] | 1
| 1 | 2 | < 10 | [12 empty] | 6
| 1 | 3 | < 10 | [12 empty] | 16
| 1 | 4 | < 10 | [12 empty] | < 10
| 1 | 5 | < 10 | [12 empty] | 12
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 471 (Indicative Value < 10 ppm) | [14 empty] | Indicative
In | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 10 | [12 empty] | 1
| 1 | 2 | < 10 | [12 empty] | 7
| 1 | 3 | < 10 | [12 empty] | 16
| 1 | 4 | < 10 | [12 empty] | < 10
| 1 | 5 | < 10 | [12 empty] | 13
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ir in OREAS 471 (Indicative Value < 10 ppm) | [14 empty] | Indicative
Ir | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 10 | [12 empty] | 1
| 1 | 2 | < 10 | [12 empty] | 8
| 1 | 3 | < 10 | [12 empty] | 16
| 1 | 4 | < 10 | [12 empty] | < 10
| 1 | 5 | < 10 | [12 empty] | 14
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 471 (Certified Value 3.23 wt.%) | [14 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 3.22 | 3.288 | 3.156 | 3.26 | 3.2426 | 3.21 | 3.1718 | 3.25 | 3.2627 | 3.1385 | 3.24 | 3.31 | 3.999 | 1
| 1 | 2 | 3.27 | 3.264 | 3.177 | 3.27 | 3.2599 | 3.21 | 3.1667 | 3.25 | 3.2756 | 3.1149 | 3.25 | 3.3 | 4.276 | #N/A
| 1 | 3 | 3.25 | 3.267 | 3.202 | 3.26 | 3.241 | 3.23 | 3.1668 | 3.22 | 3.2954 | 3.1388 | 3.26 | 3.31 | 4.349 | 16
| 1 | 4 | 3.24 | 3.27 | 3.192 | 3.28 | 3.2441 | 3.23 | 3.1725 | 3.19 | 3.2773 | 3.1237 | 3.25 | 3.32 | 4.373 | 3.2331
| 1 | 5 | 3.24 | 3.267 | 3.217 | 3.24 | 3.2507 | 3.18 | 3.1625 | 3.26 | 3.2912 | 3.1296 | 3.24 | 3.32 | 4.325 | 73
| 1 | 6 | 3.26 | 3.283 | 3.207 | 3.28 | 3.2375 | 3.21 | 3.161 | 3.22 | 3.27 | 3.1381 | 3.23 | 3.3 | 4.276
| Mean |  | 3.2467 | 3.2732 | 3.1918 | 3.265 | 3.246 | 3.2117 | 3.1669 | 3.2317 | 3.2787 | 3.1306 | 3.245 | 3.31 | 4.2663
| Median |  | 3.245 | 3.2685 | 3.197 | 3.265 | 3.2433 | 3.21 | 3.1667 | 3.235 | 3.2765 | 3.1338 | 3.245 | 3.31 | 4.3005
| Std Dev. |  | 0.017512 | 0.0098675 | 0.02223 | 0.015166 | 0.0081067 | 0.018348 | 0.0046609 | 0.026394 | 0.012452 | 0.0098082 | 0.010488 | 0.0089443 | 0.13661
| Rel.Std.Dev. |  | 0.5394% | 0.3015% | 0.6965% | 0.4645% | 0.2497% | 0.5713% | 0.1472% | 0.8167% | 0.3798% | 0.3133% | 0.3232% | 0.2702% | 3.202%
| PDM3 |  | 0.4197% | 1.239% | -1.276% | 0.9867% | 0.3984% | -0.6629% | -2.048% | -0.04427% | 1.411% | -3.17% | 0.3681% | 2.379% | 31.96%
| Z-Score (Absolute) |  | 0.01 | 0.56 | 1.12 | 0.39 | 0 | 0.71 | 1.63 | 0.29 | 0.67 | 2.38 | 0.02 | 1.32 | 21.02
| Analytical results for La2O3 in OREAS 471 (Certified Value 145 ppm) | [14 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 03 | 05 | 06 | 12 | 13 | 15 | 16 | 17 | 22 | 23 | 24 | 27 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 200 | 187.2 | 200 | 200 | < 100 | 100 | 100 | 118.4339 | 126.0339 | 100 | 104.9868 | < 12 |  | 1
| 1 | 2 | < 200 | 176.8 | 200 | 200 | < 100 | 200 | 100 | 113.1684 | 124.9293 | < 100 | 104.9868 | < 12 |  | #N/A
| 1 | 3 | < 200 | 145.6 | 200 | 200 | < 100 | 200 | 100 | 126.2647 | 121.2786 | < 100 | 106.265 | < 12 |  | 16
| 1 | 4 | < 200 | 182 | 300 | 200 | < 100 | 200 | 100 | 113.2847 | 127.3878 | < 100 | 104.7772 | < 12 |  | 145.4764
| 1 | 5 | < 200 | 187.2 | 200 | 200 | < 100 | 100 | 200 | 101.3908 | 125.2671 | < 100 | 104.9968 | < 12 |  | 74
| 1 | 6 | < 200 | 187.2 | 200 | 200 | < 100 | 100 | 200 | 100.1009 | 122.7177 | < 100 | 104.9769 | < 12
| Mean |  |  | 177.6667 | 216.6667 | 200 |  | 150 | 133.3333 | 112.1072 | 124.6024 | 100 | 105.1649
| Median |  |  | 184.6 | 200 | 200 |  | 150 | 100 | 113.2265 | 125.0982 | 100 | 104.9868
| Std Dev. |  |  | 16.2509 | 40.8248 | 0 |  | 54.7723 | 51.6398 | 10.0197 | 2.2341 |  | 0.54544
| Rel.Std.Dev. |  |  | 9.147% | 18.84% | 0% |  | 36.51% | 38.73% | 8.938% | 1.793% |  | 0.5187%
| PDM3 |  |  | 22.13% | 48.94% | 37.48% |  | 3.109% | -8.347% | -22.94% | -14.35% | -31.26% | -27.71%
| Z-Score (Absolute) |  | 0.27 | 0.88 | 1.46 | 1.21 | 1.01 | 0.47 | 0.22 | 0.09 | 0.09 | 0.89 | 0.2 | 1.67
| Analytical results for Lu2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Lu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 24 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.30496 | < 11 | [6 empty] | 1
| 1 | 2 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.29697 | < 11 | [6 empty] | #N/A
| 1 | 3 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.30396 | < 11 | [6 empty] | 16
| 1 | 4 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.32143 | < 11 | [6 empty] | < 100
| 1 | 5 | < 500 | < 100 | < 100 | < 100 | 100 | 0.31404 | < 11 | [6 empty] | 75
| 1 | 6 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.31963 | < 11
| Mean | [5 empty] | 100 | 0.31016
| Median | [5 empty] | 100 | 0.3095
| Std Dev. | [6 empty] | 0.0097107
| Rel.Std.Dev. | [6 empty] | 3.131%
PDM3
| Z-Score (Absolute) |  | 16.18 | 0 | 0 | 0 | 0.67 | 4.02 | 3.59
| Analytical results for MgO in OREAS 471 (Certified Value 1.3 wt.%) | [14 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.28 | 1.321 | 1.25 | 1.3 | 1.2953 | 1.39 | 1.3127 | 1.35 | 1.2947 | 1.4455 | 1.3 | 1.31 | 0.813 | 1
| 1 | 2 | 1.31 | 1.301 | 1.26 | 1.3 | 1.3034 | 1.39 | 1.3065 | 1.34 | 1.2946 | 1.3912 | 1.33 | 1.32 | 0.796 | #N/A
| 1 | 3 | 1.29 | 1.298 | 1.27 | 1.29 | 1.2913 | 1.34 | 1.31 | 1.33 | 1.2945 | 1.4098 | 1.3 | 1.31 | 0.779 | 16
| 1 | 4 | 1.28 | 1.317 | 1.28 | 1.3 | 1.2923 | 1.36 | 1.3174 | 1.33 | 1.2868 | 1.4405 | 1.31 | 1.3 | 0.813 | 1.3012
| 1 | 5 | 1.28 | 1.308 | 1.27 | 1.29 | 1.2908 | 1.36 | 1.3092 | 1.35 | 1.3022 | 1.4271 | 1.3 | 1.3 | 0.813 | 76
| 1 | 6 | 1.28 | 1.313 | 1.25 | 1.3 | 1.2923 | 1.38 | 1.3077 | 1.34 | 1.2925 | 1.4467 | 1.3 | 1.32 | 0.846
| Mean |  | 1.2867 | 1.3097 | 1.2633 | 1.2967 | 1.2942 | 1.37 | 1.3106 | 1.34 | 1.2942 | 1.4268 | 1.3067 | 1.31 | 0.81
| Median |  | 1.28 | 1.3105 | 1.265 | 1.3 | 1.2923 | 1.37 | 1.3096 | 1.34 | 1.2945 | 1.4338 | 1.3 | 1.31 | 0.813
| Std Dev. |  | 0.012111 | 0.0090259 | 0.012111 | 0.005164 | 0.0047633 | 0.02 | 0.0039612 | 0.0089443 | 0.0049493 | 0.022294 | 0.012111 | 0.0089443 | 0.022271
| Rel.Std.Dev. |  | 0.9412% | 0.6892% | 0.9586% | 0.3983% | 0.368% | 1.46% | 0.3022% | 0.6675% | 0.3824% | 1.562% | 0.9268% | 0.6828% | 2.75%
| PDM3 |  | -1.117% | 0.6504% | -2.91% | -0.3487% | -0.5359% | 5.287% | 0.7212% | 2.982% | -0.5366% | 9.654% | 0.4198% | 0.676% | -37.75%
| Z-Score (Absolute) |  | 1.08 | 0.16 | 2.35 | 0.54 | 0.67 | 3.43 | 0.21 | 1.81 | 0.67 | 6.51 | 0 | 0.18 | 26.92
| Analytical results for MnO in OREAS 471 (Certified Value 0.048 wt.%) | [14 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.05 | 0.05145 | 0.043 | 0.05 | 0.046336 | 0.05 | 0.05132 | 0.05 | 0.045487 | 0.072716 | 0.047 | 0.01 | 0.039 | 1
| 1 | 2 | 0.049 | 0.049392 | 0.045 | 0.04 | 0.05028 | 0.05 | 0.05132 | 0.05 | 0.043037 | 0.073087 | 0.047 | 0.01 | 0.052 | #N/A
| 1 | 3 | 0.05 | 0.050421 | 0.048 | 0.04 | 0.05028 | 0.05 | 0.052917 | 0.05 | 0.04474 | 0.080745 | 0.047 | 0.01 | 0.052 | 16
| 1 | 4 | 0.051 | 0.049392 | 0.048 | 0.06 | 0.049294 | 0.05 | 0.052169 | 0.05 | 0.045714 | 0.078225 | 0.047 | 0.02 | 0.052 | 0.048469
| 1 | 5 | 0.051 | 0.05145 | 0.042 | 0.05 | 0.048308 | 0.05 | 0.052174 | 0.05 | 0.044106 | 0.072352 | 0.047 | 0.02 | 0.052 | 77
| 1 | 6 | 0.053 | 0.05145 | 0.042 | 0.05 | 0.049294 | 0.05 | 0.053126 | 0.05 | 0.045071 | 0.072793 | 0.047 | 0.01 | 0.052
| Mean |  | 0.050667 | 0.050592 | 0.044667 | 0.048333 | 0.048965 | 0.05 | 0.052171 | 0.05 | 0.044692 | 0.074986 | 0.047 | 0.013333 | 0.049833
| Median |  | 0.0505 | 0.050936 | 0.044 | 0.05 | 0.049294 | 0.05 | 0.052172 | 0.05 | 0.044905 | 0.07294 | 0.047 | 0.01 | 0.052
| Std Dev. |  | 0.0013663 | 0.0010117 | 0.0028048 | 0.0075277 | 0.0014843 | 7.6012e-18 | 0.00076377 | 7.6012e-18 | 0.00099036 | 0.0035823 | 7.6012e-18 | 0.005164 | 0.0053072
| Rel.Std.Dev. |  | 2.697% | 2% | 6.279% | 15.57% | 3.031% | 1.52e-14% | 1.464% | 1.52e-14% | 2.216% | 4.777% | 1.617e-14% | 38.73% | 10.65%
| PDM3 |  | 4.533% | 4.38% | -7.846% | -0.2808% | 1.023% | 3.158% | 7.637% | 3.158% | -7.793% | 54.71% | -3.032% | -72.49% | 2.814%
| Z-Score (Absolute) |  | 0.29 | 0.28 | 1.22 | NA | 0.13 | NA | 0.67 | NA | 1.21 | 6.43 | 0.75 | NA | 0
Results from laboratories 4, 6, 20 and 24 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for MoO3 in OREAS 471 (Indicative Value < 1.5 ppm) | [14 empty] | Indicative
MoO3 | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 2
2
| 1 | 1 | < 100 | < 1.5 | [11 empty] | 1
| 1 | 2 | < 100 | < 1.5 | [11 empty] | 9
| 1 | 3 | < 100 | < 1.5 | [11 empty] | 16
| 1 | 4 | < 100 | < 1.5 | [11 empty] | < 1.5
| 1 | 5 | < 100 | < 1.5 | [11 empty] | 15
| 1 | 6 | < 100 | < 1.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Na2O in OREAS 471 (Certified Value 1.22 wt.%) | [14 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 23 | 24 | 27 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 1.18 | 1.222 | 1.23 | 1.21 | 1.2235 | 1.19 | 1.3 | 1.34 | 1.1894 | 1.25 | 1.17 | < 0.013 |  | 1
| 1 | 2 | 1.19 | 1.21 | 1.24 | 1.21 | 1.289 | 1.24 | 1.2897 | 1.28 | 1.1617 | 1.25 | 1.18 | < 0.013 |  | #N/A
| 1 | 3 | 1.17 | 1.211 | 1.25 | 1.2 | 1.216 | 1.23 | 1.285 | 1.28 | 1.1516 | 1.25 | 1.17 | < 0.013 |  | 16
| 1 | 4 | 1.17 | 1.218 | 1.26 | 1.22 | 1.2375 | 1.2 | 1.2941 | 1.28 | 1.1535 | 1.25 | 1.16 | < 0.013 |  | 1.2238
| 1 | 5 | 1.2 | 1.221 | 1.26 | 1.22 | 1.228 | 1.17 | 1.2903 | 1.28 | 1.1619 | 1.24 | 1.16 | < 0.013 |  | 78
| 1 | 6 | 1.19 | 1.215 | 1.28 | 1.23 | 1.226 | 1.2 | 1.2855 | 1.37 | 1.1609 | 1.24 | 1.18 | < 0.013
| Mean |  | 1.1833 | 1.2162 | 1.2533 | 1.215 | 1.2367 | 1.205 | 1.2908 | 1.305 | 1.1632 | 1.2467 | 1.17
| Median |  | 1.185 | 1.2165 | 1.255 | 1.215 | 1.227 | 1.2 | 1.29 | 1.28 | 1.1613 | 1.25 | 1.17
| Std Dev. |  | 0.012111 | 0.0050365 | 0.017512 | 0.010488 | 0.026566 | 0.025884 | 0.0056204 | 0.039875 | 0.013601 | 0.005164 | 0.0089443
| Rel.Std.Dev. |  | 1.023% | 0.4141% | 1.397% | 0.8632% | 2.148% | 2.148% | 0.4354% | 3.056% | 1.169% | 0.4142% | 0.7645%
| PDM3 |  | -3.305% | -0.6224% | 2.415% | -0.7178% | 1.053% | -1.535% | 5.474% | 6.636% | -4.954% | 1.87% | -4.395%
| Z-Score (Absolute) |  | 0.62 | 0.01 | 0.73 | 0.01 | 0.41 | 0.2 | 1.45 | 1.72 | 1.01 | 0.6 | 0.88 | 23.29
| Analytical results for Nb2O5 in OREAS 471 (Indicative Value 79 ppm) | [14 empty] | Indicative
Nb2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 06 | 21 | 23 | 27 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 50 | < 500 | 157.2 | < 100 | 20 | [8 empty] | 1
| 1 | 2 | < 50 | < 500 | 132.8 | < 100 | 21.5 | [8 empty] | 10
| 1 | 3 | < 50 | < 500 | 171.95 | < 100 | 22.9 | [8 empty] | 16
| 1 | 4 | < 50 | < 500 | 143.8 | < 100 | 22.9 | [8 empty] | 78.8403
| 1 | 5 | < 50 | < 500 | 159.4 | < 100 | 21.5 | [8 empty] | 16
| 1 | 6 | < 50 | < 500 | 173.8 | 100 | 21.5
| Mean |  |  |  | 156.4917 | 100 | 21.7167
| Median |  |  |  | 158.3 | 100 | 21.5
| Std Dev. |  |  |  | 15.9366 |  | 1.0852
| Rel.Std.Dev. |  |  |  | 10.18% |  | 4.997%
| PDM3 |  |  |  | 98.49% | 26.84% | -72.45%
| Z-Score (Absolute) |  | 0.61 | 3.53 | 1.81 | 0 | 0.67
| Analytical results for Nd2O3 in OREAS 471 (Indicative Value 84 ppm) | [14 empty] | Indicative
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 03 | 05 | 06 | 12 | 13 | 15 | 16 | 22 | 23 | 24 | 27 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | 69 | 100 | 200 | < 100 | 100 | 100 | < 100 | < 100 | 85.9886 | < 12 |  |  | 1
| 1 | 2 | < 100 | 80.5 | 200 | 200 | < 100 | 100 | 100 | < 100 | < 100 | 85.9966 | < 12 |  |  | 11
| 1 | 3 | 100 | 86.25 | 100 | 100 | < 100 | 100 | 100 | < 100 | < 100 | 86.4589 | < 12 |  |  | 16
| 1 | 4 | < 100 | 69 | < 100 | 200 | < 100 | 100 | 100 | < 100 | < 100 | 85.9007 | < 12 |  |  | 84.2166
| 1 | 5 | 100 | 92 | < 100 | 100 | < 100 | 100 | 100 | < 100 | < 100 | 85.7479 | < 12 |  |  | 17
| 1 | 6 | < 100 | 69 | 200 | 200 | < 100 | 100 | 100 | < 100 | < 100 | 85.9556 | < 12
| Mean |  | 100 | 77.625 | 150 | 166.6667 |  | 100 | 100 |  |  | 86.008
| Median |  | 100 | 74.75 | 150 | 200 |  | 100 | 100 |  |  | 85.9721
| Std Dev. |  | 0 | 10.1239 | 57.735 | 51.6398 |  | 0 | 0 |  |  | 0.23904
| Rel.Std.Dev. |  | 0% | 13.04% | 38.49% | 30.98% |  | 0% | 0% |  |  | 0.2779%
| PDM3 |  | 18.74% | -7.827% | 78.11% | 97.9% |  | 18.74% | 18.74% |  |  | 2.127%
| Z-Score (Absolute) |  | 0.27 | 0 | 0.95 | 2.17 | 0.67 | 0.55 | 0.55 | 0.67 | 0.67 | 0.2 | 1.75
| Analytical results for NiO in OREAS 471 (Indicative Value 95 ppm) | [14 empty] | Indicative
NiO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 17 | 23 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 30 | < 200 | 300 | 65 | 41.6171 | < 100 | < 13 | [6 empty] | 1
| 1 | 2 | < 30 | < 200 | 200 | 60 | 35.5077 | < 100 | < 13 | [6 empty] | 12
| 1 | 3 | < 30 | < 200 | 300 | 40 | 33.9575 | < 100 | < 13 | [6 empty] | 16
| 1 | 4 | < 30 | < 200 | 100 | 65 | 36.9396 | < 100 | < 13 | [6 empty] | 95.4911
| 1 | 5 | < 30 | < 200 | < 100 | 60 | 38.0276 | < 100 | < 13 | [6 empty] | 18
| 1 | 6 | < 30 | < 200 | 200 | 55 | 37.7901 | < 100 | < 13
| Mean |  |  |  | 220 | 57.5 | 37.3066
| Median |  |  |  | 200 | 60 | 37.3648
| Std Dev. |  |  |  | 83.666 | 9.3541 | 2.6046
| Rel.Std.Dev. |  |  |  | 38.03% | 16.27% | 6.982%
| PDM3 |  |  |  | 130.4% | -39.78% | -60.93%
| Z-Score (Absolute) |  | 0.67 | 0.96 | 2.73 | 0.14 | 0.24 | 0 | 0.84
| Analytical results for Os in OREAS 471 (Indicative Value < 10 ppm) | [14 empty] | Indicative
Os | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 10 | [12 empty] | 1
| 1 | 2 | < 10 | [12 empty] | 13
| 1 | 3 | < 10 | [12 empty] | 16
| 1 | 4 | < 10 | [12 empty] | < 10
| 1 | 5 | < 10 | [12 empty] | 19
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 471 (Certified Value 0.324 wt.%) | [14 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.32 | 0.3432 | 0.288 | 0.31 | 0.31416 | 0.32 | 0.32004 | 0.33 | 0.33803 | 0.30616 | 0.33 | 0.33 | 0.779 | 1
| 1 | 2 | 0.33 | 0.33904 | 0.292 | 0.31 | 0.31671 | 0.32 | 0.31625 | 0.33 | 0.33074 | 0.31143 | 0.33 | 0.34 | 0.848 | #N/A
| 1 | 3 | 0.32 | 0.3432 | 0.3 | 0.31 | 0.31416 | 0.32 | 0.31779 | 0.33 | 0.33516 | 0.30432 | 0.34 | 0.33 | 0.871 | 16
| 1 | 4 | 0.32 | 0.33696 | 0.292 | 0.31 | 0.31161 | 0.34 | 0.31784 | 0.33 | 0.33064 | 0.30616 | 0.33 | 0.32 | 0.871 | 0.32433
| 1 | 5 | 0.32 | 0.34216 | 0.295 | 0.31 | 0.31416 | 0.32 | 0.31816 | 0.33 | 0.33786 | 0.3057 | 0.33 | 0.32 | 0.848 | 79
| 1 | 6 | 0.33 | 0.34424 | 0.298 | 0.31 | 0.31161 | 0.34 | 0.31942 | 0.33 | 0.33585 | 0.3128 | 0.33 | 0.34 | 0.848
| Mean |  | 0.32333 | 0.34147 | 0.29417 | 0.31 | 0.31373 | 0.32667 | 0.31825 | 0.33 | 0.33471 | 0.30776 | 0.33167 | 0.33 | 0.84417
| Median |  | 0.32 | 0.34268 | 0.2935 | 0.31 | 0.31416 | 0.32 | 0.318 | 0.33 | 0.3355 | 0.30616 | 0.33 | 0.33 | 0.848
| Std Dev. |  | 0.005164 | 0.0028418 | 0.0044008 | 0 | 0.0019196 | 0.010328 | 0.0013393 | 0 | 0.0033098 | 0.0034663 | 0.0040825 | 0.0089443 | 0.033855
| Rel.Std.Dev. |  | 1.597% | 0.8322% | 1.496% | 0% | 0.6118% | 3.162% | 0.4208% | 0% | 0.9888% | 1.126% | 1.231% | 2.71% | 4.01%
| PDM3 |  | -0.3063% | 5.285% | -9.299% | -4.417% | -3.266% | 0.7215% | -1.874% | 1.749% | 3.203% | -5.107% | 2.263% | 1.749% | 160.3%
| Z-Score (Absolute) |  | 0.27 | 1.19 | 2.6 | 1.34 | 1.04 | 0 | 0.67 | 0.27 | 0.64 | 1.51 | 0.4 | 0.27 | 41.45
| Analytical results for PbO in OREAS 471 (Indicative Value 56 ppm) | [14 empty] | Indicative
PbO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 17 | 23 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 200 | < 100 | 100 | < 100 | 40.9518 | < 100 | 40.9 | [6 empty] | 1
| 1 | 2 | < 200 | < 100 | 80 | < 100 | 33.1077 | < 100 | 44.2 | [6 empty] | 14
| 1 | 3 | < 200 | < 100 | 80 | < 100 | 36.516 | < 100 | 45.2 | [6 empty] | 16
| 1 | 4 | < 200 | < 100 | 100 | < 100 | 35.411 | < 100 | 47.4 | [6 empty] | 56.039
| 1 | 5 | < 200 | < 100 | 80 | < 100 | 39.27 | < 100 | 44.2 | [6 empty] | 20
| 1 | 6 | < 200 | < 100 | 80 | < 100 | 37.3763 | < 100 | 44.2
| Mean |  |  |  | 86.6667 |  | 37.1055 |  | 44.35
| Median |  |  |  | 80 |  | 36.9462 |  | 44.2
| Std Dev. |  |  |  | 10.328 |  | 2.784 |  | 2.0974
| Rel.Std.Dev. |  |  |  | 11.92% |  | 7.503% |  | 4.729%
| PDM3 |  |  |  | 54.65% |  | -33.79% |  | -20.86%
| Z-Score (Absolute) |  | 5.96 | 0 | 4.37 | 0 | 1.54 | 0 | 0.67
| Analytical results for Pd in OREAS 471 (Indicative Value < 1000 ppb) | [14 empty] | Indicative
Pd | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppb
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 1000 | [12 empty] | 1
| 1 | 2 | < 1000 | [12 empty] | 15
| 1 | 3 | < 1000 | [12 empty] | 16
| 1 | 4 | < 1000 | [12 empty] | < 1000
| 1 | 5 | < 1000 | [12 empty] | 21
| 1 | 6 | < 1000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pm in OREAS 471 (Indicative Value 118 ppm) | [14 empty] | Indicative
Pm | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | 110 | [12 empty] | 1
| 1 | 2 | 116 | [12 empty] | 16
| 1 | 3 | 120 | [12 empty] | 16
| 1 | 4 | 126 | [12 empty] | 117.5
| 1 | 5 | 118 | [12 empty] | 22
| 1 | 6 | 115
| Mean |  | 117.5
| Median |  | 117
| Std Dev. |  | 5.3572
| Rel.Std.Dev. |  | 4.559%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr6O11 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 06 | 12 | 13 | 15 | 16 | 22 | 23 | 24 | 27 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | 27.3731 | < 24 |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | 100 | < 100 | 100 | 26.76 | < 24 |  |  |  | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | 100 | < 100 | < 100 | 26.8478 | < 24 |  |  |  | 16
| 1 | 4 | < 100 | 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | 27.0995 | < 24 |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | 27.0286 | < 24 |  |  |  | 80
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | < 100 | 27.3661 | < 24
| Mean |  |  | 100 |  |  | 100 | 100 |  | 100 | 27.0792
| Median |  |  | 100 |  |  | 100 | 100 |  | 100 | 27.064
| Std Dev. | [5 empty] | 0 | 0 |  |  | 0.25576
| Rel.Std.Dev. | [5 empty] | 0% | 0% |  |  | 0.9445%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 0 | 0 | 1.35 | 3.37 | 0 | 0.67 | 1.85 | 3.07
| Analytical results for Pt in OREAS 471 (Indicative Value < 1000 ppb) | [14 empty] | Indicative
Pt | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppb
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 1000 | [12 empty] | 1
| 1 | 2 | < 1000 | [12 empty] | 17
| 1 | 3 | < 1000 | [12 empty] | 16
| 1 | 4 | < 1000 | [12 empty] | < 1000
| 1 | 5 | < 1000 | [12 empty] | 23
| 1 | 6 | < 1000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ra in OREAS 471 (Indicative Value < 20 ppm) | [14 empty] | Indicative
Ra | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 20 | [12 empty] | 1
| 1 | 2 | < 20 | [12 empty] | 18
| 1 | 3 | < 20 | [12 empty] | 16
| 1 | 4 | < 20 | [12 empty] | < 20
| 1 | 5 | < 20 | [12 empty] | 24
| 1 | 6 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb2O in OREAS 471 (Indicative Value 119 ppm) | [14 empty] | Indicative
Rb2O | Round | Replicate | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 23 | 27 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 100 | < 100 | 197.9 | [10 empty] | 1
| 1 | 2 | 100 | < 100 | 194.7 | [10 empty] | 19
| 1 | 3 | 100 | 100 | 196.8 | [10 empty] | 16
| 1 | 4 | 100 | < 100 | 205.6 | [10 empty] | 118.7583
| 1 | 5 | 100 | < 100 | 196.8 | [10 empty] | 25
| 1 | 6 | 100 | < 100 | 195.8
| Mean |  | 100 | 100 | 197.9333
| Median |  | 100 | 100 | 196.8
| Std Dev. |  | 0 |  | 3.9078
| Rel.Std.Dev. |  | 0% |  | 1.974%
| PDM3 |  | -15.8% | -15.8% | 66.67%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.59
| Analytical results for Re in OREAS 471 (Indicative Value < 10 ppm) | [14 empty] | Indicative
Re | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | < 10 | [12 empty] | 1
| 1 | 2 | < 10 | [12 empty] | 3
| 1 | 3 | < 10 | [12 empty] | 16
| 1 | 4 | < 10 | [12 empty] | < 10
| 1 | 5 | < 10 | [12 empty] | 26
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rh in OREAS 471 (Indicative Value < 1 ppm) | [14 empty] | Indicative
Rh | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 2
2
| 1 | 1 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | [12 empty] | 5
| 1 | 3 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | [12 empty] | < 1
| 1 | 5 | < 1 | [12 empty] | 11
| 1 | 6 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 471 (Indicative Value < 1 ppm) | [14 empty] | Indicative
Ru | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 2
2
| 1 | 1 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | [12 empty] | 6
| 1 | 3 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | [12 empty] | < 1
| 1 | 5 | < 1 | [12 empty] | 12
| 1 | 6 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb2O3 in OREAS 471 (Indicative Value < 36 ppm) | [14 empty] | Indicative
Sb2O3 | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 100 | < 36 | [11 empty] | 1
| 1 | 2 | < 100 | < 36 | [11 empty] | 7
| 1 | 3 | < 100 | < 36 | [11 empty] | 16
| 1 | 4 | < 100 | < 36 | [11 empty] | < 36
| 1 | 5 | < 100 | < 36 | [11 empty] | 13
| 1 | 6 | < 100 | < 36
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Sc2O3 in OREAS 471 (Indicative Value 16.5 ppm) | [14 empty] | Indicative
Sc2O3 | Round | Replicate | Lab | Lab | Lab | [10 empty] | 1
| No | No | 23 | 24 | 27 | [10 empty] | ppm
|  |  | BF*XRF | PPP*XRF | BF*XRF | [10 empty] | 1
1
| 1 | 1 | < 100 | 16.6667 | < 46 | [10 empty] | 1
| 1 | 2 | < 100 | 16.6587 | < 46 | [10 empty] | 8
| 1 | 3 | < 100 | 16.1564 | < 46 | [10 empty] | 16
| 1 | 4 | < 100 | 16.6437 | < 46 | [10 empty] | 16.4954
| 1 | 5 | < 100 | 16.6128 | < 46 | [10 empty] | 14
| 1 | 6 | < 100 | 16.2343 | < 46
| Mean |  |  | 16.4954
| Median |  |  | 16.6282
| Std Dev. |  |  | 0.23445
| Rel.Std.Dev. |  |  | 1.421%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | 2.8 | 0.67 | 0
| Analytical results for Se in OREAS 471 (Indicative Value < 1 ppm) | [14 empty] | Indicative
Se | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 2
2
| 1 | 1 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | [12 empty] | 1
| 1 | 3 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | [12 empty] | < 1
| 1 | 5 | < 1 | [12 empty] | 15
| 1 | 6 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 471 (Certified Value 63.35 wt.%) | [14 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 63.08 | 63.535 | 60.43 | 62.43 | 64.1573 | 64.3 | 63.2767 | 64.66 | 64.2177 | 63.8569 | 64.01 | 62.55 | 57.459 | 1
| 1 | 2 | 63.48 | 63.273 | 61.1 | 62.05 | 64.3947 | 62.3 | 63.148 | 64.13 | 64.2353 | 63.3639 | 64.01 | 62.57 | 60.368 | #N/A
| 1 | 3 | 63.43 | 63.322 | 61.47 | 61.97 | 64.0354 | 63.9 | 63.1578 | 63.94 | 64.0462 | 63.1208 | 64.02 | 62.56 | 60.176 | 16
| 1 | 4 | 63.24 | 63.277 | 61.75 | 61.74 | 64.1183 | 64.1 | 63.2857 | 63.95 | 64.2036 | 63.878 | 63.94 | 62.55 | 59.769 | 63.3514
| 1 | 5 | 63.11 | 63.299 | 61.9 | 61.49 | 64.1877 | 64.5 | 63.1808 | 64.33 | 64.1762 | 64.128 | 64.08 | 62.53 | 60.112 | 81
| 1 | 6 | 63.52 | 63.603 | 61.71 | 62.07 | 64.0399 | 64.7 | 63.244 | 64.21 | 64.2986 | 63.7157 | 64.02 | 62.54 | 60.154
| Mean |  | 63.31 | 63.3848 | 61.3933 | 61.9583 | 64.1556 | 63.9667 | 63.2155 | 64.2033 | 64.1963 | 63.6772 | 64.0133 | 62.55 | 59.673
| Median |  | 63.335 | 63.3105 | 61.59 | 62.01 | 64.1378 | 64.2 | 63.2124 | 64.17 | 64.2107 | 63.7863 | 64.015 | 62.55 | 60.133
| Std Dev. |  | 0.19246 | 0.14533 | 0.54848 | 0.31953 | 0.13222 | 0.8641 | 0.060967 | 0.26964 | 0.084188 | 0.37015 | 0.044572 | 0.014142 | 1.1019
| Rel.Std.Dev. |  | 0.304% | 0.2293% | 0.8934% | 0.5157% | 0.2061% | 1.351% | 0.09644% | 0.42% | 0.1311% | 0.5813% | 0.06963% | 0.02261% | 1.847%
| PDM3 |  | -0.06538% | 0.05274% | -3.091% | -2.199% | 1.269% | 0.9712% | -0.2146% | 1.345% | 1.334% | 0.5143% | 1.045% | -1.265% | -5.806%
| Z-Score (Absolute) |  | 0.07 | 0 | 1.74 | 1.25 | 0.67 | 0.51 | 0.15 | 0.72 | 0.71 | 0.26 | 0.55 | 0.73 | 3.25
| Analytical results for Sm2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 06 | 12 | 13 | 15 | 16 | 22 | 23 | 24 | 27 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 13.3945 | < 58 |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 13.1778 | < 58 |  |  |  | #N/A
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | < 100 | < 100 | 100 | 12.9751 | < 58 |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 12.9651 | < 58 |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 13.0769 | < 58 |  |  |  | 82
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 13.2307 | < 58
| Mean | [5 empty] | 100 |  |  | 100 | 13.1367
| Median | [5 empty] | 100 |  |  | 100 | 13.1274
| Std Dev. | [9 empty] | 0.16499
| Rel.Std.Dev. | [9 empty] | 1.256%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for SnO2 in OREAS 471 (Indicative Value < 100 ppm) | [14 empty] | Indicative
SnO2 | Round | Replicate | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 21 | 23 | 27 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 200 | < 100 | < 100 | 68.6 | [9 empty] | 1
| 1 | 2 | < 200 | < 100 | < 100 | 69.8 | [9 empty] | 10
| 1 | 3 | < 200 | < 100 | < 100 | 69.8 | [9 empty] | 16
| 1 | 4 | < 200 | < 100 | < 100 | 73.6 | [9 empty] | < 100
| 1 | 5 | < 200 | < 100 | < 100 | 72.4 | [9 empty] | 16
| 1 | 6 | < 200 | < 100 | < 100 | 73.6
| Mean |  |  |  |  | 71.3
| Median |  |  |  |  | 71.1
| Std Dev. |  |  |  |  | 2.1716
| Rel.Std.Dev. |  |  |  |  | 3.046%
PDM3
| Z-Score (Absolute) |  | 2.49 | 0.67 | 0.67 | 0.67
| Analytical results for SO3 in OREAS 471 (Certified Value 0.068 wt.%) | [14 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 20 | 23 | 24 | 27 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.066 | 0.06 | < 0.05 | 0.06825 | 0.043 | 0.07 | 0.07 | < 0.025 | [5 empty] | 1
| 1 | 2 | 0.066 | 0.06 | < 0.05 | 0.07875 | 0.052 | 0.08 | 0.06 | < 0.025 | [5 empty] | #N/A
| 1 | 3 | 0.067 | 0.07 | < 0.05 | 0.06615 | 0.048 | 0.07 | 0.07 | < 0.025 | [5 empty] | 16
| 1 | 4 | 0.066 | 0.06 | < 0.05 | 0.07035 | 0.041 | 0.07 | 0.08 | < 0.025 | [5 empty] | 0.067967
| 1 | 5 | 0.067 | 0.06 | < 0.05 | 0.07035 | 0.049 | 0.07 | 0.08 | < 0.025 | [5 empty] | 83
| 1 | 6 | 0.067 | 0.06 | < 0.05 | 0.06615 | 0.044 | 0.07 | 0.06 | < 0.025
| Mean |  | 0.0665 | 0.061667 |  | 0.07 | 0.046167 | 0.071667 | 0.07
| Median |  | 0.0665 | 0.06 |  | 0.0693 | 0.046 | 0.07 | 0.07
| Std Dev. |  | 0.00054772 | 0.0040825 |  | 0.0046801 | 0.0041673 | 0.0040825 | 0.0089443
| Rel.Std.Dev. |  | 0.8236% | 6.62% |  | 6.686% | 9.027% | 5.696% | 12.78%
| PDM3 |  | -2.158% | -9.269% |  | 2.992% | -32.07% | 5.444% | 2.992%
| Z-Score (Absolute) |  | 0.24 | 0.24 | 3.9 | 0.59 | 1.79 | 0.76 | 0.59 | 5.15
| Analytical results for SrO in OREAS 471 (Indicative Value 541 ppm) | [14 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 17 | 20 | 23 | 24 | 27 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 350 | 540 | 500 | 530.4 | 513.6389 | 520 | 600 | 300 | 741.5 |  |  |  |  | 1
| 1 | 2 | 340 | 560 | 600 | 586.5 | 511.2681 | 520 | 600 | 300 | 729.7 |  |  |  |  | 11
| 1 | 3 | 360 | 570 | 400 | 576.3 | 514.0706 | 532 | 500 | 300 | 745 |  |  |  |  | 16
| 1 | 4 | 330 | 540 | 500 | 622.2 | 511.2976 | 544 | 600 | 200 | 772.2 |  |  |  |  | 540.9681
| 1 | 5 | 330 | 550 | 600 | 530.4 | 512.0522 | 532 | 500 | 200 | 739.1 |  |  |  |  | 17
| 1 | 6 | 350 | 560 | 500 | 525.3 | 512.0555 | 520 | 500 | 300 | 735.6
| Mean |  | 343.3333 | 553.3333 | 516.6667 | 561.85 | 512.3971 | 528 | 550 | 266.6667 | 743.85
| Median |  | 345 | 555 | 500 | 553.35 | 512.0539 | 526 | 550 | 300 | 740.3
| Std Dev. |  | 12.1106 | 12.1106 | 75.2773 | 39.4275 | 1.1884 | 9.798 | 54.7723 | 51.6398 | 14.8419
| Rel.Std.Dev. |  | 3.527% | 2.189% | 14.57% | 7.017% | 0.2319% | 1.856% | 9.959% | 19.36% | 1.995%
| PDM3 |  | -36.53% | 2.286% | -4.492% | 3.86% | -5.281% | -2.397% | 1.67% | -50.71% | 37.5%
| Z-Score (Absolute) |  | 4.97 | 0.67 | 0.31 | 0.9 | 0.43 | 0 | 0.58 | 7.03 | 5.79
| Analytical results for Ta2O5 in OREAS 471 (Indicative Value < 100 ppm) | [14 empty] | Indicative
Ta2O5 | Round | Replicate | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 06 | 23 | 27 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | 5 | < 500 | < 100 | < 12 | [9 empty] | 1
| 1 | 2 | 10 | < 500 | < 100 | < 12 | [9 empty] | 12
| 1 | 3 | 10 | < 500 | < 100 | < 12 | [9 empty] | 16
| 1 | 4 | 10 | < 500 | < 100 | < 12 | [9 empty] | < 100
| 1 | 5 | 10 | < 500 | < 100 | < 12 | [9 empty] | 18
| 1 | 6 | 10 | < 500 | < 100 | < 12
| Mean |  | 9.1667
| Median |  | 10
| Std Dev. |  | 2.0412
| Rel.Std.Dev. |  | 22.27%
PDM3
| Z-Score (Absolute) |  | 0.63 | 6.77 | 0.63 | 0.72
| Analytical results for Tb4O7 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Tb4O7 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 22 | 24 | 27 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 500 | < 100 | < 100 | < 100 | < 100 | < 100 | 1.3068 | 1247.9 | [5 empty] | 1
| 1 | 2 | < 500 | < 100 | < 100 | < 100 | < 100 | < 100 | 1.3084 | 1270.3 | [5 empty] | #N/A
| 1 | 3 | < 500 | < 100 | < 100 | < 100 | < 100 | < 100 | 1.3004 | 1290.3 | [5 empty] | 16
| 1 | 4 | < 500 | < 100 | < 100 | < 100 | < 100 | < 100 | 1.2803 | 1350.3 | [5 empty] | < 100
| 1 | 5 | < 500 | < 100 | < 100 | < 100 | < 100 | < 100 | 1.251 | 1270.3 | [5 empty] | 84
| 1 | 6 | < 500 | < 100 | < 100 | < 100 | < 100 | < 100 | 1.3245 | 1270.3
| Mean | [7 empty] | 1.2952 | 1283.2333
| Median | [7 empty] | 1.3036 | 1270.3
| Std Dev. | [7 empty] | 0.025958 | 35.4912
| Rel.Std.Dev. | [7 empty] | 2.004% | 2.766%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tc in OREAS 471 (Indicative Value < 1 ppm) | [14 empty] | Indicative
Tc | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 2
2
| 1 | 1 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | [12 empty] | 13
| 1 | 3 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | [12 empty] | < 1
| 1 | 5 | < 1 | [12 empty] | 19
| 1 | 6 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 471 (Indicative Value 40.7 ppm) | [14 empty] | Indicative
Te | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 1
1
| 1 | 1 | 39 | [12 empty] | 1
| 1 | 2 | 36 | [12 empty] | 2
| 1 | 3 | 40 | [12 empty] | 16
| 1 | 4 | 44 | [12 empty] | 40.6667
| 1 | 5 | 40 | [12 empty] | 20
| 1 | 6 | 45
| Mean |  | 40.6667
| Median |  | 40
| Std Dev. |  | 3.3267
| Rel.Std.Dev. |  | 8.18%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 471 (Indicative Value < 50 ppm) | [14 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 03 | 05 | 06 | 12 | 13 | 15 | 16 | 21 | 22 | 23 | 24 | 27 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 44 | 18.1035 | < 90 | < 50 | < 50 | < 50 | < 50 | < 30 | < 90 | < 90 | 23.3625 | 92 |  | 1
| 1 | 2 | < 44 | 18.1035 | < 90 | < 50 | < 50 | < 50 | < 50 | < 30 | < 90 | < 90 | 23.4029 | 99 |  | 15
| 1 | 3 | < 44 | 18.1035 | < 90 | < 50 | < 50 | < 50 | < 50 | < 30 | < 90 | < 90 | 23.2019 | 104 |  | 16
| 1 | 4 | < 44 | 18.1035 | < 90 | < 50 | < 50 | < 50 | < 50 | < 30 | < 90 | < 90 | 23.2432 | 104 |  | < 50
| 1 | 5 | < 44 | 18.1035 | < 90 | < 50 | 90 | < 50 | < 50 | < 30 | < 90 | < 90 | 23.1888 | 100 |  | 21
| 1 | 6 | < 44 | 27.1553 | < 90 | < 50 | < 50 | < 50 | < 50 | < 30 | < 90 | < 90 | 23.1265 | 102
| Mean |  |  | 19.6122 |  |  | 90 | [5 empty] | 23.2543 | 100.1667
| Median |  |  | 18.1035 |  |  | 90 | [5 empty] | 23.2226 | 101
| Std Dev. |  |  | 3.6954 | [8 empty] | 0.10704 | 4.4907
| Rel.Std.Dev. |  |  | 18.84% | [8 empty] | 0.4603% | 4.483%
PDM3
| Z-Score (Absolute) |  | 0.27 | 0.47 | 1.66 | 0 | 0.95 | 0 | 0 | 0.88 | 1.66 | 1.66 | 0.15 | 6.59
| Analytical results for TiO2 in OREAS 471 (Certified Value 0.942 wt.%) | [14 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.91 | 0.93534 | 0.97 | 0.93 | 0.92558 | 0.95 | 0.95113 | 0.95 | 0.96728 | 0.9055 | 0.95 | 0.98 | 1.168 | 1
| 1 | 2 | 0.92 | 0.94758 | 0.98 | 0.94 | 0.93293 | 0.91 | 0.94051 | 0.95 | 0.95318 | 0.8972 | 0.95 | 0.97 | 1.284 | #N/A
| 1 | 3 | 0.92 | 0.93024 | 0.98 | 0.94 | 0.92768 | 0.92 | 0.94171 | 0.93 | 0.96088 | 0.8845 | 0.95 | 0.98 | 1.301 | 16
| 1 | 4 | 0.92 | 0.9384 | 0.99 | 0.94 | 0.92453 | 0.95 | 0.94482 | 0.94 | 0.96059 | 0.8813 | 0.95 | 0.99 | 1.318 | 0.94247
| 1 | 5 | 0.91 | 0.93126 | 1 | 0.93 | 0.9303 | 0.95 | 0.94377 | 0.95 | 0.96266 | 0.8709 | 0.94 | 0.99 | 1.284 | 85
| 1 | 6 | 0.93 | 0.94248 | 0.99 | 0.94 | 0.92978 | 0.94 | 0.94614 | 0.94 | 0.95212 | 0.9074 | 0.95 | 0.97 | 1.268
| Mean |  | 0.91833 | 0.93755 | 0.985 | 0.93667 | 0.92846 | 0.93667 | 0.94468 | 0.94333 | 0.95945 | 0.89113 | 0.94833 | 0.98 | 1.2705
| Median |  | 0.92 | 0.93687 | 0.985 | 0.94 | 0.92873 | 0.945 | 0.9443 | 0.945 | 0.96073 | 0.89085 | 0.95 | 0.98 | 1.284
| Std Dev. |  | 0.0075277 | 0.006686 | 0.010488 | 0.005164 | 0.0031456 | 0.017512 | 0.003762 | 0.008165 | 0.0057965 | 0.014546 | 0.0040825 | 0.0089443 | 0.053035
| Rel.Std.Dev. |  | 0.8197% | 0.7131% | 1.065% | 0.5513% | 0.3388% | 1.87% | 0.3982% | 0.8655% | 0.6042% | 1.632% | 0.4305% | 0.9127% | 4.174%
| PDM3 |  | -2.561% | -0.5218% | 4.513% | -0.6155% | -1.486% | -0.6155% | 0.2349% | 0.09185% | 1.802% | -5.447% | 0.6224% | 3.982% | 34.81%
| Z-Score (Absolute) |  | 1.13 | 0.26 | 1.89 | 0.3 | 0.67 | 0.3 | 0.06 | 0 | 0.73 | 2.37 | 0.23 | 1.66 | 14.84
| Analytical results for Tl in OREAS 471 (Indicative Value < 50 ppm) | [14 empty] | Indicative
Tl | Round | Replicate | Lab | [12 empty] | 1
| No | No | 27 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | [12 empty] | 16
| 1 | 3 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | [12 empty] | < 50
| 1 | 5 | < 50 | [12 empty] | 22
| 1 | 6 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Tm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 24 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.33501 | < 57 | [6 empty] | 1
| 1 | 2 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.3437 | < 57 | [6 empty] | #N/A
| 1 | 3 | < 500 | < 100 | < 100 | 100 | < 100 | 0.33591 | < 57 | [6 empty] | 16
| 1 | 4 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.33002 | < 57 | [6 empty] | < 100
| 1 | 5 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.33781 | < 57 | [6 empty] | 86
| 1 | 6 | < 500 | < 100 | < 100 | < 100 | < 100 | 0.33491 | < 57
| Mean |  |  |  |  | 100 |  | 0.33623
| Median |  |  |  |  | 100 |  | 0.33546
| Std Dev. | [6 empty] | 0.0044761
| Rel.Std.Dev. | [6 empty] | 1.331%
PDM3
| Z-Score (Absolute) |  | 16.18 | 0 | 0 | 0.67 | 0 | 4.02 | 1.74
| Analytical results for U3O8 in OREAS 471 (Indicative Value < 100 ppm) | [14 empty] | Indicative
U3O8 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 05 | 06 | 21 | 23 | 24 | 27 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 10 | < 100 | < 24 | < 100 | 2.2267 | < 12 | [7 empty] | 1
| 1 | 2 |  | < 100 | < 24 | < 100 | 2.2584 | < 12 | [7 empty] | 17
| 1 | 3 |  | < 100 | < 24 | < 100 | 2.2651 | < 12 | [7 empty] | 16
| 1 | 4 | 5 | < 100 | < 24 | < 100 | 2.2865 | < 12 | [7 empty] | < 100
| 1 | 5 |  | < 100 | < 24 | < 100 | 2.2588 | < 12 | [7 empty] | 23
| 1 | 6 | 10 | < 100 | < 24 | < 100 | 2.2445 | < 12
| Mean |  | 8.3333 |  |  |  | 2.2567
| Median |  | 10 |  |  |  | 2.2586
| Std Dev. |  | 2.8868 |  |  |  | 0.020074
| Rel.Std.Dev. |  | 34.64% |  |  |  | 0.8896%
PDM3
| Z-Score (Absolute) |  | 0.19 | 4.5 | 0.19 | 4.5 | 0.88 | 0.47
| Analytical results for V2O5 in OREAS 471 (Certified Value 175 ppm) | [14 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 17 | 21 | 23 | 27 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 170 | < 200 | 160 | 100 | 180 | 159.2653 | 253.5 | 200 | 192.8 |  |  |  |  | 1
| 1 | 2 | 160 | < 200 | 200 | 100 | 190 | 124.797 | 251.1 | 200 | 224.9 |  |  |  |  | #N/A
| 1 | 3 | 200 | < 200 | 160 | 100 | 160 | 151.0295 | 243 | 200 | 223.2 |  |  |  |  | 16
| 1 | 4 | 190 | < 200 | 140 | 200 | 190 | 155.9079 | 215.2 | 200 | 217.8 |  |  |  |  | 174.597
| 1 | 5 | 200 | < 200 | 180 | 200 | 170 | 151.5278 | 260.6 | 200 | 203.5 |  |  |  |  | 87
| 1 | 6 | 160 | < 200 | 180 | 100 | 160 | 142.2778 | 244.7 | 200 | 208.9
| Mean |  | 180 |  | 170 | 133.3333 | 175 | 147.4675 | 244.6833 | 200 | 211.85
| Median |  | 180 |  | 170 | 100 | 175 | 151.2786 | 247.9 | 200 | 213.35
| Std Dev. |  | 18.9737 |  | 20.9762 | 51.6398 | 13.784 | 12.4939 | 15.7756 | 0 | 12.4568
| Rel.Std.Dev. |  | 10.54% |  | 12.34% | 38.73% | 7.877% | 8.472% | 6.447% | 0% | 5.88%
| PDM3 |  | 3.095% |  | -2.633% | -23.63% | 0.2308% | -15.54% | 40.14% | 14.55% | 21.34%
| Z-Score (Absolute) |  | 0.12 | 1.84 | 0.12 | 1.02 | 0 | 0.67 | 1.71 | 0.61 | 0.9
| Analytical results for WO3 in OREAS 471 (Indicative Value < 100 ppm) | [14 empty] | Indicative
WO3 | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 23 | 27 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | < 130 | [11 empty] | 1
| 1 | 2 | < 100 | < 130 | [11 empty] | 18
| 1 | 3 | < 100 | < 130 | [11 empty] | 16
| 1 | 4 | < 100 | < 130 | [11 empty] | < 100
| 1 | 5 | < 100 | < 130 | [11 empty] | 24
| 1 | 6 | < 100 | < 130
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Y in OREAS 471 (Indicative Value 62 ppm) | [14 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 06 | 12 | 13 | 15 | 16 | 22 | 23 | 24 | 27 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 39 | < 80 | 100 | < 100 | 100 | 100 | < 80 | < 80 | 27.4196 | 34 |  |  |  | 1
| 1 | 2 | < 39 | < 80 | 100 | < 100 | 100 | 100 | < 80 | < 80 | 27.2498 | 38 |  |  |  | 19
| 1 | 3 | < 39 | < 80 | < 100 | < 100 | 100 | 100 | < 80 | < 80 | 27.275 | 39 |  |  |  | 16
| 1 | 4 | 39 | < 80 | 100 | < 100 | 100 | 100 | < 80 | < 80 | 26.7788 | 40 |  |  |  | 62.0712
| 1 | 5 | < 39 | < 80 | < 100 | < 100 | 100 | 100 | < 80 | < 80 | 26.9825 | 39 |  |  |  | 25
| 1 | 6 | < 39 | < 80 | 100 | < 100 | 100 | 100 | < 80 | < 80 | 27.0516 | 39
| Mean |  | 39 |  | 100 |  | 100 | 100 |  |  | 27.1262 | 38.1667
| Median |  | 39 |  | 100 |  | 100 | 100 |  |  | 27.1507 | 39
| Std Dev. |  |  |  | 0 |  | 0 | 0 |  |  | 0.23256 | 2.137
| Rel.Std.Dev. |  |  |  | 0% |  | 0% | 0% |  |  | 0.8573% | 5.599%
| PDM3 |  | -37.17% |  | 61.11% |  | 61.11% | 61.11% |  |  | -56.3% | -38.51%
| Z-Score (Absolute) |  | 0.97 | 0 | 2.59 | 0.63 | 3.57 | 3.57 | 0 | 0 | 0.72 | 0.07
| Analytical results for Yb2O3 in OREAS 471 (Certified Value < 100 ppm) | [14 empty] | Certified
Yb2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 06 | 12 | 13 | 15 | 16 | 24 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | 2.2264 | < 34 | [6 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | 2.1804 | < 34 | [6 empty] | #N/A
| 1 | 3 | < 100 | 100 | < 100 | < 100 | < 100 | 2.2311 | < 34 | [6 empty] | 16
| 1 | 4 | < 100 | 100 | < 100 | < 100 | < 100 | 2.2214 | < 34 | [6 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | 2.2081 | < 34 | [6 empty] | 88
| 1 | 6 | < 100 | 100 | < 100 | < 100 | < 100 | 2.2628 | < 34
| Mean |  |  | 100 |  |  |  | 2.2217
| Median |  |  | 100 |  |  |  | 2.2239
| Std Dev. |  |  | 0 |  |  |  | 0.027165
| Rel.Std.Dev. |  |  | 0% |  |  |  | 1.223%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for ZnO in OREAS 471 (Certified Value 100 ppm) | [14 empty] | Certified
ZnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 17 | 23 | 27 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 200 | 110 | 100 | 109.2 | 84.4986 | < 100 | 113.3 | [6 empty] | 1
| 1 | 2 | < 200 | 110 | 100 | 114.4 | 83.354 | 100 | 125.7 | [6 empty] | #N/A
| 1 | 3 | < 200 | 110 | 100 | 109.2 | 84.4775 | 100 | 128.2 | [6 empty] | 16
| 1 | 4 | < 200 | 120 | 100 | 98.8 | 84.1039 | < 100 | 134.4 | [6 empty] | 99.8418
| 1 | 5 | < 200 | 110 | 100 | 104 | 83.7584 | < 100 | 129.5 | [6 empty] | 89
| 1 | 6 | < 200 | 130 | 100 | 109.2 | 83.7191 | < 100 | 124.5
| Mean |  |  | 115 | 100 | 107.4667 | 83.9853 | 100 | 125.9333
| Median |  |  | 110 | 100 | 109.2 | 83.9312 | 100 | 126.95
| Std Dev. |  |  | 8.3666 | 0 | 5.3705 | 0.45623 | 0 | 7.0888
| Rel.Std.Dev. |  |  | 7.275% | 0% | 4.997% | 0.5432% | 0% | 5.629%
| PDM3 |  |  | 15.18% | 0.1584% | 7.637% | -15.88% | 0.1584% | 26.13%
| Z-Score (Absolute) |  | 0 | 0.67 | 0 | 0.34 | 0.72 | 1.5 | 1.17
| Analytical results for ZrO2 in OREAS 471 (Certified Value 403 ppm) | [14 empty] | Certified
ZrO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 17 | 21 | 23 | 24 | 27 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 450 | 800 | 400 | 385 | 418.6935 | 315 | 400 | 400 | 578.1 |  |  |  |  | 1
| 1 | 2 | 410 | 700 | 400 | 400 | 417.5358 | 363.1 | 400 | 300 | 571.4 |  |  |  |  | #N/A
| 1 | 3 | 420 | 800 | 400 | 425 | 415.7761 | 385 | 400 | 400 | 584.9 |  |  |  |  | 16
| 1 | 4 | 440 | 700 | 400 | 385 | 419.2407 | 343.8 | 400 | 500 | 603.8 |  |  |  |  | 402.6519
| 1 | 5 | 400 | 800 | 400 | 430 | 415.2776 | 347.4 | 400 | 500 | 584.9 |  |  |  |  | 90
| 1 | 6 | 390 | 700 | 500 | 380 | 414.8562 | 440.7 | 400 | 300 | 578.1
| Mean |  | 418.3333 | 750 | 416.6667 | 400.8333 | 416.8966 | 365.8333 | 400 | 400 | 583.5333
| Median |  | 415 | 750 | 400 | 392.5 | 416.6559 | 355.25 | 400 | 400 | 581.5
| Std Dev. |  | 23.1661 | 54.7723 | 40.8248 | 21.7754 | 1.8532 | 43.3339 | 0 | 89.4427 | 11.1421
| Rel.Std.Dev. |  | 5.538% | 7.303% | 9.798% | 5.433% | 0.4445% | 11.85% | 0% | 22.36% | 1.909%
| PDM3 |  | 3.895% | 86.27% | 3.481% | -0.4516% | 3.538% | -9.144% | -0.6586% | -0.6586% | 44.92%
| Z-Score (Absolute) |  | 0.07 | 13.49 | 0 | 0.64 | 0.01 | 2.06 | 0.67 | 0.67 | 6.75
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C630: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C648: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C666: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C684: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C702: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C720: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C738: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C756: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C774: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C792: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C810: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C828: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C846: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C864: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C882: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1026: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1116: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1188: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1206: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1224: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1242: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1260: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1278: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1296: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1314: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1332: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 471 (Certified Value 4.73 wt.%) | [18 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 12 | 13 | 15 | 16 | 17 | 20 | 21 | 22 | 23 | 24 | 27 | wt.%
|  |  | BF*OES/MS | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | BF*XRF | BF*OES/MS | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | PF*OES/MS | 2
2
| 1 | 1 | 5.32 | 4.5909 | 4.41 | 5.2 | 5.4583 | 4.38 | 4.23 | 4.37 | 4.2 | 4.21 | 5.3128 | 4.5 | 4.593 | 4.21 | 4.33 | 5.85 | 5.08 | 1
| 1 | 2 | 5.32 | 4.6127 | 4.15 | 5.2 | 5.378 | 5.39 | 4.25 | 4.41 | 4.2 | 4.17 | 5.3143 | 4.7 | 4.4898 | 4.17 | 4.31 | 5.86 | 5.16 | #N/A
| 1 | 3 | 5.34 | 4.6342 | 4.28 | 5.19 | 5.2316 | 4.38 | 4.17 | 4.39 | 4.18 | 4.17 | 5.3364 | 4.6 | 4.6899 | 4.25 | 4.34 | 5.85 | 5.13 | 16
| 1 | 4 | 5.3 | 4.5949 | 4.3 | 5.2 | 5.4901 | 4.31 | 4.2 | 4.42 | 4.18 | 4.15 | 5.3192 | 4.8 | 4.5554 | 4.25 | 4.42 | 5.84 | 5.24 | 4.7259
| 1 | 5 | 5.31 | 4.624 | 4.07 | 5.21 | 5.5855 | 4.16 | 4.34 | 4.35 | 4.2 | 4.22 | 5.3101 | 4.5 | 4.5622 | 4.25 | 4.33 | 5.84 | 5.17 | 92
| 1 | 6 | 5.3 | 4.6131 | 4.21 | 5.23 | 5.5003 | 4.19 | 4.19 | 4.38 | 4.2 | 4.15 | 5.3012 | 4.5 | 4.5529 | 4.2 | 4.37 | 5.86 | 5.2
| Mean |  | 5.315 | 4.6116 | 4.2367 | 5.205 | 5.4406 | 4.4683 | 4.23 | 4.3867 | 4.1933 | 4.1783 | 5.3157 | 4.6 | 4.5739 | 4.2217 | 4.35 | 5.85 | 5.1633
| Median |  | 5.315 | 4.6129 | 4.245 | 5.2 | 5.4742 | 4.345 | 4.215 | 4.385 | 4.2 | 4.17 | 5.3136 | 4.55 | 4.5588 | 4.23 | 4.335 | 5.85 | 5.165
| Std Dev. |  | 0.015166 | 0.016585 | 0.11994 | 0.013784 | 0.12241 | 0.46102 | 0.060992 | 0.02582 | 0.010328 | 0.029944 | 0.01177 | 0.12649 | 0.066069 | 0.033714 | 0.039497 | 0.0089443 | 0.055377
| Rel.Std.Dev. |  | 0.2853% | 0.3596% | 2.831% | 0.2648% | 2.25% | 10.32% | 1.442% | 0.5886% | 0.2463% | 0.7167% | 0.2214% | 2.75% | 1.444% | 0.7986% | 0.908% | 0.1529% | 1.073%
| PDM3 |  | 12.47% | -2.418% | -10.35% | 10.14% | 15.12% | -5.45% | -10.49% | -7.178% | -11.27% | -11.59% | 12.48% | -2.664% | -3.217% | -10.67% | -7.954% | 23.79% | 9.256%
| Z-Score (Absolute) |  | 1.42 | 0.07 | 0.65 | 1.21 | 1.66 | 0.2 | 0.66 | 0.36 | 0.73 | 0.76 | 1.42 | 0.05 | 0 | 0.67 | 0.43 | 2.44 | 1.13
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 471 (Indicative Value 0.042 wt.%) |  |  | Indicative
C | Round | Replicate | Lab | 1
| No | No | 20 | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.04 | 22
| 1 | 3 | 0.04 | 16
| 1 | 4 | 0.04 | 0.041667
| 1 | 5 | 0.04 | 28
| 1 | 6 | 0.04
| Mean |  | 0.041667
| Median |  | 0.04
| Std Dev. |  | 0.0040825
| Rel.Std.Dev. |  | 9.798%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 471 (Indicative Value 0.03 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 20 | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.03 | 22
| 1 | 3 | 0.02 | 16
| 1 | 4 | < 0.02 | 0.03
| 1 | 5 | 0.03 | 28
| 1 | 6 | 0.05
| Mean |  | 0.034
| Median |  | 0.03
| Std Dev. |  | 0.011402
| Rel.Std.Dev. |  | 33.53%
| PDM3 |  | 13.33%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 471 (Indicative Value 0.049 ppm) | [27 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 25 | 27 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 3
3
| 1 | 1 | 0.6 | < 0.1 | < 0.3 | 0.05 | 0.04 | 0.19 | < 0.1 | < 0.05 | < 0.05 | 0.02 | 0.03 | 0.01 | 0.05 | 0.03 | < 0.05 | 0.14 | < 0.02 | 0.31 | 0.2 | 0.11 | [6 empty] | 1
| 1 | 2 | 0.6 | < 0.1 | < 0.3 | 0.06 | 0.04 | 0.2 | < 0.1 | < 0.05 | < 0.05 | 0.02 | 0.03 | 0.01 | 0.04 | 0.02 | < 0.05 | 0.135 | < 0.02 | 0.315 | 0.3 | 0.1 | [6 empty] | 1
| 1 | 3 | 0.7 | < 0.1 | < 0.3 | 0.05 | 0.04 | 0.21 | < 0.1 | < 0.05 | < 0.05 | 0.02 | 0.03 | 0.03 | 0.04 | 0.02 | < 0.05 | 0.103 | < 0.02 | 0.305 | 0.3 | 0.09 | [6 empty] | 16
| 1 | 4 | 0.7 | < 0.1 | < 0.3 | 0.05 | 0.03 | 0.21 | < 0.1 | < 0.05 | < 0.05 | 0.03 | 0.03 | 0.01 | 0.04 | 0.03 | < 0.05 | 0.102 | < 0.02 | 0.316 | 0.2 | 0.1 | [6 empty] | 0.049
| 1 | 5 | 0.6 | < 0.1 | < 0.3 | 0.05 | 0.03 | 0.2 | < 0.1 | < 0.05 | < 0.05 | 0.03 | 0.02 | 0.03 | 0.03 | 0.02 | < 0.05 | 0.111 | < 0.02 | 0.301 | 0.2 | 0.09 | [6 empty] | 30
| 1 | 6 | 0.6 | < 0.1 | < 0.3 | 0.05 | 0.03 | 0.2 | < 0.1 | < 0.05 | < 0.05 | 0.03 | 0.03 | 0.02 | 0.04 | 0.02 | < 0.05 | 0.125 | < 0.02 | 0.283 | 0.2 | 0.11
| Mean |  | 0.63333 |  |  | 0.051667 | 0.035 | 0.20167 |  |  |  | 0.025 | 0.028333 | 0.018333 | 0.04 | 0.023333 |  | 0.11933 |  | 0.305 | 0.23333 | 0.1
| Median |  | 0.6 |  |  | 0.05 | 0.035 | 0.2 |  |  |  | 0.025 | 0.03 | 0.015 | 0.04 | 0.02 |  | 0.118 |  | 0.3075 | 0.2 | 0.1
| Std Dev. |  | 0.05164 |  |  | 0.0040825 | 0.0054772 | 0.0075277 |  |  |  | 0.0054772 | 0.0040825 | 0.0098319 | 0.0063246 | 0.005164 |  | 0.016379 |  | 0.012215 | 0.05164 | 0.0089443
| Rel.Std.Dev. |  | 8.154% |  |  | 7.902% | 15.65% | 3.733% |  |  |  | 21.91% | 14.41% | 53.63% | 15.81% | 22.13% |  | 13.73% |  | 4.005% | 22.13% | 8.944%
| PDM3 |  | 1193% |  |  | 5.442% | -28.57% | 311.6% |  |  |  | -48.98% | -42.18% | -62.59% | -18.37% | -52.38% |  | 143.5% |  | 522.4% | 376.2% | 104.1%
| Z-Score (Absolute) |  | 19.04 | 0.16 | 3.4 | 0.22 | 0.32 | 5.07 | 0.16 | 0.65 | 0.65 | 0.65 | 0.54 | 0.86 | 0.16 | 0.7 | 0.65 | 2.41 | 1.13 | 8.42 | 6.1 | 1.78
| Analytical results for Al in OREAS 471 (Certified Value 8.46 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 8.57 | 8.2554 | 8.3745 | 8.93 | 6.21 | 8.41 | 8.28 | 8.3844 | 7.87 | 8.7777 | 8.84 | 7.82 | 8.75 | 8.36 | 8.59 | 8.5728 | 7.14 | 8.43 | 8.1654 | 8.7304 | 8.42 | 0.0016299 | 8.58 |  |  |  | 1
| 1 | 2 | 7.93 | 8.3545 | 8.2833 | 8.87 | 7.32 | 8.55 | 8.29 | 8.6301 | 8.01 | 8.3281 | 8.58 | 8.06 | 8.34 | 8.62 | 8.65 | 8.4733 | 7.25 | 8.5 | 8.3134 | 8.7256 | 8.58 | 0.0016489 | 8.71 |  |  |  | #N/A
| 1 | 3 | 8.36 | 8.3296 | 8.2294 | 8.88 | 7.39 | 8.37 | 8.06 | 8.262 | 8.06 | 8.4646 | 8.49 | 8.44 | 8.43 | 8.75 | 8.32 | 8.5432 | 6.7 | 8.47 | 8.6377 | 8.7158 | 8.58 | 0.0016365 | 8.6 |  |  |  | 16
| 1 | 4 | 8.47 | 8.309 | 8.3096 | 9.11 | 6.52 | 8.37 | 8.15 | 8.6336 | 8.23 | 8.5589 | 8.88 | 8.4 | 8.63 | 8.26 | 8.28 | 8.6293 | 7.19 | 8.65 | 8.4513 | 8.7167 | 8.55 | 0.0015427 | 8.7 |  |  |  | 8.4609
| 1 | 5 | 7.99 | 8.4609 | 8.1885 | 8.86 | 7.31 | 8.64 | 8.11 | 8.4576 | 8.16 | 8.4667 | 8.64 | 8.32 | 8.47 | 8.44 | 8 | 8.5489 | 7.14 | 8.47 | 8.3044 | 8.741 | 8.46 | 0.0016037 | 8.68 |  |  |  | 94
| 1 | 6 | 8.69 | 8.5505 | 8.3625 | 8.75 | 6.83 | 8.21 | 8.16 | 8.5024 | 7.39 | 8.5743 | 8.94 | 8.37 | 8.64 | 8.24 | 8.42 | 8.6886 | 6.91 | 8.43 | 7.8094 | 8.7291 | 8.43 | 0.0015873 | 8.65
| Mean |  | 8.335 | 8.3766 | 8.2913 | 8.9 | 6.93 | 8.425 | 8.175 | 8.4784 | 7.9533 | 8.5284 | 8.7283 | 8.235 | 8.5433 | 8.445 | 8.3767 | 8.576 | 7.055 | 8.4917 | 8.2803 | 8.7264 | 8.5033 | 0.0016082 | 8.6533
| Median |  | 8.415 | 8.342 | 8.2965 | 8.875 | 7.07 | 8.39 | 8.155 | 8.48 | 8.035 | 8.5128 | 8.74 | 8.345 | 8.55 | 8.4 | 8.37 | 8.5608 | 7.14 | 8.47 | 8.3089 | 8.7273 | 8.505 | 0.0016168 | 8.665
| Std Dev. |  | 0.31085 | 0.10891 | 0.073198 | 0.11866 | 0.49084 | 0.15123 | 0.09225 | 0.14406 | 0.3027 | 0.15046 | 0.18269 | 0.24395 | 0.15436 | 0.20354 | 0.23517 | 0.074683 | 0.20869 | 0.082077 | 0.28056 | 0.0094239 | 0.075011 | 3.9178e-05 | 0.053541
| Rel.Std.Dev. |  | 3.729% | 1.3% | 0.8828% | 1.333% | 7.083% | 1.795% | 1.128% | 1.699% | 3.806% | 1.764% | 2.093% | 2.962% | 1.807% | 2.41% | 2.807% | 0.8708% | 2.958% | 0.9666% | 3.388% | 0.108% | 0.8821% | 2.436% | 0.6187%
| PDM3 |  | -1.488% | -0.9959% | -2.005% | 5.19% | -18.09% | -0.4243% | -3.379% | 0.2062% | -5.999% | 0.7976% | 3.161% | -2.67% | 0.9743% | -0.1879% | -0.9956% | 1.361% | -16.62% | 0.3636% | -2.135% | 3.138% | 0.5015% | -99.98% | 2.274%
| Z-Score (Absolute) |  | 0.47 | 0.28 | 0.68 | 2.2 | 7.13 | 0.05 | 1.23 | 0.21 | 2.28 | 0.44 | 1.39 | 0.95 | 0.51 | 0.05 | 0.28 | 0.67 | 6.54 | 0.27 | 0.73 | 1.38 | 0.32 | NA | 1.03
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for As in OREAS 471 (Certified Value 5.26 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 25 | 27 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 7.3 | < 10 | 8.8 | 6.1 | 5 | 9.7 | 5 | 5 | 11 | 4.6 | 4.2 | 4 | 4.6 | 3.6 | 4.1 | < 10 | 7.9 | 5.1 | 22.552 | 5.7 | 5.8 | [5 empty] | 1
| 1 | 2 | 7.1 | < 10 | 8.8 | 6.4 | 5 | 9.8 | 5 | 4.7 | 12 | 4.5 | 5.1 | 4.6 | 3.8 | 4 | 4.4 | < 10 | 8 | 5.2 | 23.473 | 5.7 | 6.2 | [5 empty] | 2
| 1 | 3 | 7 | < 10 | 9.1 | 6 | 5 | 9.9 | 5 | 4.7 | 10 | 4.6 | 4.4 | 4 | 4.5 | 4 | 4.2 | < 10 | 7.2 | 4.9 | 24.121 | 5.7 | 6.4 | [5 empty] | 16
| 1 | 4 | 7 | < 10 | 9.2 | 6.3 | 4 | 11.4 | 5 | 4.9 | 17 | 4.8 | 4.8 | 4.2 | 4.6 | 4 | 4.3 | < 10 | 7.4 | 5 | 25.56 | 5.5 | 5.9 | [5 empty] | 5.2644
| 1 | 5 | 7.3 | < 10 | 9 | 6.2 | 5 | 10.5 | 5 | 4.9 | 11 | 4.7 | 4.3 | 4.2 | 4.1 | 4.6 | 4.1 | < 10 | 7.2 | 5 | 20.443 | 5.7 | 6.3 | [5 empty] | 95
| 1 | 6 | 7.3 | < 10 | 9.2 | 5.9 | 5 | 10.4 | 5 | 5 | 15 | 4.7 | 4.4 | 4.2 | 4.4 | 4 | 4.3 | < 10 | 7.8 | 5 | 25.305 | 5.8 | 5.9
| Mean |  | 7.1667 |  | 9.0167 | 6.15 | 4.8333 | 10.2833 | 5 | 4.8667 | 12.6667 | 4.65 | 4.5333 | 4.2 | 4.3333 | 4.0333 | 4.2333 |  | 7.5833 | 5.0333 | 23.5757 | 5.6833 | 6.0833
| Median |  | 7.2 |  | 9.05 | 6.15 | 5 | 10.15 | 5 | 4.9 | 11.5 | 4.65 | 4.4 | 4.2 | 4.45 | 4 | 4.25 |  | 7.6 | 5 | 23.797 | 5.7 | 6.05
| Std Dev. |  | 0.15055 |  | 0.18348 | 0.18708 | 0.40825 | 0.63692 | 0 | 0.13663 | 2.7325 | 0.10488 | 0.34448 | 0.21909 | 0.32042 | 0.32042 | 0.12111 |  | 0.36009 | 0.10328 | 1.9025 | 0.098319 | 0.24833
| Rel.Std.Dev. |  | 2.101% |  | 2.035% | 3.042% | 8.447% | 6.194% | 0% | 2.807% | 21.57% | 2.256% | 7.599% | 5.216% | 7.394% | 7.944% | 2.861% |  | 4.748% | 2.052% | 8.07% | 1.73% | 4.082%
| PDM3 |  | 36.14% |  | 71.28% | 16.82% | -8.188% | 95.34% | -5.022% | -7.554% | 140.6% | -11.67% | -13.89% | -20.22% | -17.69% | -23.38% | -19.59% |  | 44.05% | -4.388% | 347.8% | 7.959% | 15.56%
| Z-Score (Absolute) |  | 1.81 | 0.01 | 3.37 | 0.96 | NA | 4.44 | NA | 0.13 | NA | 0.31 | 0.41 | 0.69 | 0.58 | 0.83 | 0.66 | 0.01 | 2.16 | 0.01 | 15.64 | 0.56 | 0.9
Results from laboratories 6, 8 and 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Ba in OREAS 471 (Certified Value 0.247 wt.%) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.253 | 0.23557 | 0.2538 | 0.273 | 0.221 | 0.2512 | 0.251 | 0.25726 | 0.2227 | 0.24805 | 0.26 | 0.235 | 0.254 | 0.244 | 0.25 | 0.23739 | 0.2484 | 0.2318 | 0.2581 | 0.23667 | 0.25545 | 0.25332 | 0.2393 |  |  |  | 1
| 1 | 2 | 0.254 | 0.23881 | 0.256 | 0.276 | 0.2407 | 0.2516 | 0.248 | 0.25147 | 0.227 | 0.24177 | 0.254 | 0.241 | 0.239 | 0.253 | 0.255 | 0.23682 | 0.2456 | 0.2364 | 0.2563 | 0.23889 | 0.25556 | 0.25739 | 0.2415 |  |  |  | #N/A
| 1 | 3 | 0.255 | 0.23837 | 0.251 | 0.271 | 0.2242 | 0.2519 | 0.253 | 0.24877 | 0.2284 | 0.24548 | 0.247 | 0.246 | 0.245 | 0.253 | 0.242 | 0.24819 | 0.2476 | 0.2316 | 0.2556 | 0.24566 | 0.25359 | 0.25692 | 0.2403 |  |  |  | 16
| 1 | 4 | 0.261 | 0.23817 | 0.2545 | 0.273 | 0.2198 | 0.2523 | 0.248 | 0.25704 | 0.2337 | 0.24774 | 0.259 | 0.256 | 0.248 | 0.241 | 0.245 | 0.23563 | 0.2495 | 0.2365 | 0.2577 | 0.2395 | 0.25629 | 0.25606 | 0.242 |  |  |  | 0.24747
| 1 | 5 | 0.259 | 0.24082 | 0.2519 | 0.282 | 0.23 | 0.2531 | 0.257 | 0.2551 | 0.233 | 0.24731 | 0.251 | 0.242 | 0.245 | 0.247 | 0.238 | 0.23414 | 0.247 | 0.236 | 0.2542 | 0.23955 | 0.25861 | 0.25523 | 0.2426 |  |  |  | 96
| 1 | 6 | 0.261 | 0.24419 | 0.2546 | 0.281 | 0.2291 | 0.2519 | 0.247 | 0.26017 | 0.2243 | 0.25182 | 0.262 | 0.24 | 0.25 | 0.247 | 0.248 | 0.22931 | 0.2511 | 0.2322 | 0.2575 | 0.23146 | 0.25648 | 0.25401 | 0.2416
| Mean |  | 0.25717 | 0.23932 | 0.25363 | 0.276 | 0.22747 | 0.252 | 0.25067 | 0.25497 | 0.22818 | 0.24703 | 0.2555 | 0.24333 | 0.24683 | 0.2475 | 0.24633 | 0.23691 | 0.2482 | 0.23408 | 0.25657 | 0.23862 | 0.256 | 0.25549 | 0.24122
| Median |  | 0.257 | 0.23859 | 0.25415 | 0.2745 | 0.22665 | 0.2519 | 0.2495 | 0.25607 | 0.2277 | 0.24752 | 0.2565 | 0.2415 | 0.2465 | 0.247 | 0.2465 | 0.23622 | 0.248 | 0.2341 | 0.2569 | 0.2392 | 0.25593 | 0.25565 | 0.24155
| Std Dev. |  | 0.0036009 | 0.0029158 | 0.0018576 | 0.0045607 | 0.0076873 | 0.00065115 | 0.0038297 | 0.0041779 | 0.0044772 | 0.0033057 | 0.0058224 | 0.0071461 | 0.0051153 | 0.0048062 | 0.0060222 | 0.0062373 | 0.0019339 | 0.0024417 | 0.0014909 | 0.0046113 | 0.0016375 | 0.0016083 | 0.0012057
| Rel.Std.Dev. |  | 1.4% | 1.218% | 0.7324% | 1.652% | 3.38% | 0.2584% | 1.528% | 1.639% | 1.962% | 1.338% | 2.279% | 2.937% | 2.072% | 1.942% | 2.445% | 2.633% | 0.7792% | 1.043% | 0.5811% | 1.932% | 0.6396% | 0.6295% | 0.4998%
| PDM3 |  | 3.918% | -3.293% | 2.49% | 11.53% | -8.083% | 1.83% | 1.291% | 3.03% | -7.794% | -0.1787% | 3.245% | -1.672% | -0.2575% | 0.01188% | -0.4596% | -4.266% | 0.2947% | -5.41% | 3.676% | -3.575% | 3.446% | 3.24% | -2.527%
| Z-Score (Absolute) |  | 0.82 | 0.69 | 0.52 | 2.41 | 1.69 | 0.38 | 0.27 | 0.63 | 1.63 | 0.04 | 0.68 | 0.35 | 0.06 | 0 | 0.1 | 0.89 | 0.06 | 1.13 | 0.77 | 0.75 | 0.72 | 0.67 | 0.53
| Analytical results for Be in OREAS 471 (Certified Value 1.99 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.1 | 2.075 | 1.83 | 2.43 | 2.1 | 2.02 | 1.9 | 2.18 | 2.1 | 2.2 | 2 | 1.89 | 1.96 | 1.91 | 1.95 | 2.0179 | 1.91 | 2 | 2 | 2.189 | 0.99124 | 2 | 2.2559 | 2.09 |  |  | 1
| 1 | 2 | 1.9 | 2.1098 | 1.85 | 2.29 | 2.2 | 2.06 | 1.9 | 2.05 | 2.17 | 2.18 | 1.96 | 1.95 | 1.88 | 1.97 | 1.98 | 1.9191 | 1.97 | 2 | 2 | 2.362 | 0.9802 | 2 | 1.8193 | 2.04 |  |  | #N/A
| 1 | 3 | 1.9 | 1.9125 | 1.81 | 2.32 | 2.1 | 1.95 | 1.9 | 2.01 | 2.11 | 2.09 | 1.92 | 1.98 | 1.91 | 1.96 | 1.88 | 2.0067 | 1.91 | 2 | 2 | 2.557 | 0.96817 | 2.1 | 2.0864 | 2.11 |  |  | 16
| 1 | 4 | 2 | 2.0393 | 1.82 | 2.47 | 1.8 | 2.13 | 2 | 2.04 | 2.11 | 2.19 | 2.01 | 2.1 | 1.92 | 1.9 | 1.9 | 1.9196 | 1.86 | 2 | 2 | 2.424 | 0.96681 | 2 | 1.9433 | 2.02 |  |  | 1.9904
| 1 | 5 | 1.8 | 2.0323 | 1.82 | 2.26 | 1.8 | 2.09 | 1.9 | 2.18 | 2.14 | 2 | 1.94 | 1.97 | 1.91 | 1.93 | 1.87 | 1.9126 | 1.89 | 2 | 2 | 2.42 | 0.9557 | 2 | 1.834 | 2.07 |  |  | 97
| 1 | 6 | 2.1 | 1.9899 | 1.84 | 2.15 | 1.9 | 2.05 | 1.9 | 2.04 | 2.04 | 2.09 | 2.03 | 1.94 | 1.93 | 1.97 | 1.93 | 1.9512 | 2 | 2 | 2 | 2.11 | 0.97892 | 2 | 2.2802 | 2.09
| Mean |  | 1.9667 | 2.0265 | 1.8283 | 2.32 | 1.9833 | 2.05 | 1.9167 | 2.0833 | 2.1117 | 2.125 | 1.9767 | 1.9717 | 1.9183 | 1.94 | 1.9183 | 1.9545 | 1.9233 | 2 | 2 | 2.3437 | 0.97351 | 2.0167 | 2.0365 | 2.07
| Median |  | 1.95 | 2.0358 | 1.825 | 2.305 | 2 | 2.055 | 1.9 | 2.045 | 2.11 | 2.135 | 1.98 | 1.96 | 1.915 | 1.945 | 1.915 | 1.9354 | 1.91 | 2 | 2 | 2.391 | 0.97355 | 2 | 2.0149 | 2.08
| Std Dev. |  | 0.12111 | 0.06905 | 0.01472 | 0.11662 | 0.17224 | 0.061644 | 0.040825 | 0.07607 | 0.043551 | 0.078677 | 0.043205 | 0.070261 | 0.026394 | 0.030984 | 0.042622 | 0.046867 | 0.052026 | 0 | 0 | 0.16532 | 0.01248 | 0.040825 | 0.20339 | 0.034059
| Rel.Std.Dev. |  | 6.158% | 3.407% | 0.8051% | 5.027% | 8.684% | 3.007% | 2.13% | 3.651% | 2.062% | 3.702% | 2.186% | 3.564% | 1.376% | 1.597% | 2.222% | 2.398% | 2.705% | 0% | 0% | 7.054% | 1.282% | 2.024% | 9.987% | 1.645%
| PDM3 |  | -1.192% | 1.813% | -8.142% | 16.56% | -0.3548% | 2.995% | -3.704% | 4.669% | 6.093% | 6.763% | -0.6898% | -0.941% | -3.621% | -2.532% | -3.621% | -1.803% | -3.369% | 0.4825% | 0.4825% | 17.75% | -51.09% | 1.32% | 2.318% | 3.999%
| Z-Score (Absolute) |  | 0.15 | 0.51 | 1.66 | 3.72 | 0.04 | 0.77 | 0.69 | 1.13 | 1.44 | 1.59 | 0.04 | 0.09 | 0.67 | 0.44 | 0.67 | 0.28 | 0.62 | NA | NA | 3.98 | 11.01 | 0.4 | 0.62 | 0.98
Results from laboratories 19 and 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 471 (Certified Value 0.3 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.31 | 0.27385 | 0.27 | 0.31 | 0.28 | 2.07 | 0.29 | 0.32 | 0.4 | 0.32 | 0.3 | 0.31 | 0.3 | 0.29 | 0.3 | 0.31542 | 0.32 | 0.28 | 0.32 | 0.299 | < 0.09 | 0.24 | 0.20074 | 0.3 |  |  | 1
| 1 | 2 | 0.3 | 0.28559 | 0.26 | 0.31 | 0.29 | 2.1 | 0.31 | 0.31 | 0.33 | 0.32 | 0.3 | 0.34 | 0.3 | 0.3 | 0.33 | 0.31844 | 0.33 | 0.28 | 0.31 | 0.306 | < 0.09 | 0.25 | 0.21822 | 0.29 |  |  | #N/A
| 1 | 3 | 0.31 | 0.28176 | 0.25 | 0.32 | 0.28 | 1.99 | 0.28 | 0.32 | 0.49 | 0.33 | 0.3 | 0.33 | 0.29 | 0.3 | 0.3 | 0.32144 | 0.32 | 0.26 | 0.31 | 0.306 | < 0.09 | 0.26 | 0.23809 | 0.3 |  |  | 16
| 1 | 4 | 0.3 | 0.28757 | 0.26 | 0.31 | 0.27 | 1.97 | 0.3 | 0.31 | 0.35 | 0.31 | 0.31 | 0.33 | 0.3 | 0.29 | 0.3 | 0.30901 | 0.33 | 0.27 | 0.31 | 0.303 | < 0.09 | 0.25 | 0.2286 | 0.29 |  |  | 0.30067
| 1 | 5 | 0.3 | 0.28061 | 0.24 | 0.3 | 0.28 | 2.09 | 0.3 | 0.33 | 0.36 | 0.31 | 0.3 | 0.32 | 0.3 | 0.3 | 0.29 | 0.31119 | 0.32 | 0.28 | 0.31 | 0.31 | < 0.09 | 0.24 | 0.22861 | 0.3 |  |  | 98
| 1 | 6 | 0.31 | 0.28775 | 0.25 | 0.3 | 0.27 | 2.16 | 0.31 | 0.32 | 0.4 | 0.33 | 0.29 | 0.32 | 0.29 | 0.29 | 0.3 | 0.31121 | 0.33 | 0.27 | 0.32 | 0.3 | < 0.09 | 0.24 | 0.22227 | 0.28
| Mean |  | 0.305 | 0.28286 | 0.255 | 0.30833 | 0.27833 | 2.0633 | 0.29833 | 0.31833 | 0.38833 | 0.32 | 0.3 | 0.325 | 0.29667 | 0.295 | 0.30333 | 0.31445 | 0.325 | 0.27333 | 0.31333 | 0.304 |  | 0.24667 | 0.22276 | 0.29333
| Median |  | 0.305 | 0.28368 | 0.255 | 0.31 | 0.28 | 2.08 | 0.3 | 0.32 | 0.38 | 0.32 | 0.3 | 0.325 | 0.3 | 0.295 | 0.3 | 0.31332 | 0.325 | 0.275 | 0.31 | 0.3045 |  | 0.245 | 0.22544 | 0.295
| Std Dev. |  | 0.0054772 | 0.0053108 | 0.010488 | 0.0075277 | 0.0075277 | 0.071461 | 0.01169 | 0.0075277 | 0.057067 | 0.0089443 | 0.0063246 | 0.010488 | 0.005164 | 0.0054772 | 0.013663 | 0.0048258 | 0.0054772 | 0.008165 | 0.005164 | 0.0041473 |  | 0.008165 | 0.012721 | 0.008165
| Rel.Std.Dev. |  | 1.796% | 1.878% | 4.113% | 2.441% | 2.705% | 3.463% | 3.919% | 2.365% | 14.7% | 2.795% | 2.108% | 3.227% | 1.741% | 1.857% | 4.504% | 1.535% | 1.685% | 2.987% | 1.648% | 1.364% |  | 3.31% | 5.711% | 2.784%
| PDM3 |  | 1.442% | -5.923% | -15.19% | 2.55% | -7.428% | 586.3% | -0.7756% | 5.876% | 29.16% | 6.431% | -0.2213% | 8.094% | -1.33% | -1.884% | 0.8873% | 4.585% | 8.094% | -9.091% | 4.213% | 1.109% |  | -17.96% | -25.91% | -2.439%
| Z-Score (Absolute) |  | 0.13 | 0.72 | 1.8 | 0.26 | 0.9 | 67.88 | 0.13 | 0.64 | 3.34 | 0.71 | 0.06 | 0.9 | 0.19 | 0.26 | 0.06 | 0.49 | 0.9 | 1.09 | 0.45 | 0.09 | 9.9 | 2.12 | 3.04 | 0.32
| Analytical results for Ca in OREAS 471 (Certified Value 0.862 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 1 | 0.84787 | 0.9005 | 0.87 | 0.68 | 0.81 | 0.851 | 0.8672 | 0.843 | 0.8763 | 0.87 | 0.83 | 0.88 | 0.82 | 0.87 | 0.8675 | 0.91 | 0.87 | 0.90489 | 0.88805 | 0.86 | 0.0012856 | 0.88 |  |  |  | 1
| 1 | 2 | 0.95 | 0.85078 | 0.9132 | 0.87 | 0.8 | 0.84 | 0.843 | 0.8789 | 0.855 | 0.8552 | 0.86 | 0.85 | 0.82 | 0.86 | 0.89 | 0.8744 | 0.83 | 0.87 | 1.0007 | 0.88778 | 0.86 | 0.0013125 | 0.89 |  |  |  | #N/A
| 1 | 3 | 0.97 | 0.85964 | 0.8983 | 0.87 | 0.76 | 0.83 | 0.839 | 0.8548 | 0.863 | 0.8841 | 0.83 | 0.86 | 0.85 | 0.86 | 0.84 | 0.8619 | 0.77 | 0.87 | 0.94116 | 0.88674 | 0.87 | 0.0013049 | 0.89 |  |  |  | 16
| 1 | 4 | 1.01 | 0.83173 | 0.9103 | 0.88 | 0.7 | 0.84 | 0.841 | 0.8971 | 0.874 | 0.8867 | 0.87 | 0.9 | 0.87 | 0.82 | 0.84 | 0.8737 | 0.81 | 0.89 | 0.94674 | 0.88581 | 0.86 | 0.0012932 | 0.89 |  |  |  | 0.86171
| 1 | 5 | 0.94 | 0.855 | 0.899 | 0.87 | 0.74 | 0.8 | 0.843 | 0.8729 | 0.876 | 0.8785 | 0.84 | 0.85 | 0.85 | 0.83 | 0.81 | 0.8652 | 0.76 | 0.87 | 0.92442 | 0.88701 | 0.85 | 0.0012861 | 0.89 |  |  |  | 99
| 1 | 6 | 1 | 0.85611 | 0.9071 | 0.86 | 0.69 | 0.81 | 0.838 | 0.8806 | 0.832 | 0.8928 | 0.88 | 0.85 | 0.87 | 0.83 | 0.85 | 0.872 | 0.8 | 0.87 | 0.88722 | 0.88667 | 0.85 | 0.0013038 | 0.89
| Mean |  | 0.97833 | 0.85019 | 0.90473 | 0.87 | 0.72833 | 0.82167 | 0.8425 | 0.87525 | 0.85717 | 0.87893 | 0.85833 | 0.85667 | 0.85667 | 0.83667 | 0.85 | 0.86912 | 0.81333 | 0.87333 | 0.93419 | 0.88701 | 0.85833 | 0.0012977 | 0.88833
| Median |  | 0.985 | 0.85289 | 0.9038 | 0.87 | 0.72 | 0.82 | 0.842 | 0.8759 | 0.859 | 0.8813 | 0.865 | 0.85 | 0.86 | 0.83 | 0.845 | 0.86975 | 0.805 | 0.87 | 0.93279 | 0.88688 | 0.86 | 0.0012985 | 0.89
| Std Dev. |  | 0.029269 | 0.0099401 | 0.0063317 | 0.0063246 | 0.046655 | 0.017224 | 0.0046368 | 0.014198 | 0.017383 | 0.013036 | 0.019408 | 0.023381 | 0.021602 | 0.018619 | 0.027568 | 0.0050452 | 0.053914 | 0.008165 | 0.039464 | 0.00081277 | 0.0075277 | 1.104e-05 | 0.0040825
| Rel.Std.Dev. |  | 2.992% | 1.169% | 0.6998% | 0.727% | 6.406% | 2.096% | 0.5504% | 1.622% | 2.028% | 1.483% | 2.261% | 2.729% | 2.522% | 2.225% | 3.243% | 0.5805% | 6.629% | 0.9349% | 4.224% | 0.09163% | 0.877% | 0.8507% | 0.4596%
| PDM3 |  | 13.53% | -1.337% | 4.993% | 0.9626% | -15.48% | -4.646% | -2.229% | 1.572% | -0.5267% | 1.999% | -0.3913% | -0.5847% | -0.5847% | -2.906% | -1.358% | 0.8601% | -5.614% | 1.349% | 8.411% | 2.937% | -0.3913% | -99.85% | 3.09%
| Z-Score (Absolute) |  | 4.94 | 0.34 | 1.91 | 0.48 | 5.35 | 1.51 | 0.65 | 0.7 | 0.05 | 0.85 | 0 | 0.07 | 0.07 | 0.89 | 0.34 | 0.44 | 1.85 | 0.62 | 3.12 | 1.18 | 0 | NA | 1.24
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Cd in OREAS 471 (Indicative Value 0.028 ppm) | [27 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | < 0.3 | < 0.5 | 0.05 | 0.04 | 0.03 | 0.07 | < 0.05 | < 0.02 | 0.04 | < 0.02 | 0.02 | 0.02 | 0.02 | < 0.02 | 0.02 | 0.02764 | < 0.02 | < 0.02 | < 0.02 | < 0.1 | 0.12 | 0.14 |  |  |  |  | 1
| 1 | 2 | < 0.3 | < 0.5 | 0.05 | 0.04 | 0.03 | 0.07 | < 0.05 | < 0.02 | 0.03 | 0.03 | 0.02 | < 0.02 | < 0.02 | 0.02 | < 0.02 | 0.02762 | 0.03 | < 0.02 | < 0.02 | < 0.1 | 0.13 | 0.13 |  |  |  |  | 8
| 1 | 3 | < 0.3 | < 0.5 | 0.05 | 0.05 | 0.03 | 0.08 | < 0.05 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | < 0.02 | 0.02 | < 0.02 | 0.03238 | 0.03 | < 0.02 | < 0.02 | < 0.1 | 0.12 | 0.14 |  |  |  |  | 16
| 1 | 4 | < 0.3 | < 0.5 | 0.05 | 0.04 | < 0.02 | 0.07 | < 0.05 | < 0.02 | 0.04 | 0.02 | < 0.02 | 0.02 | 0.02 | < 0.02 | < 0.02 | 0.02874 | < 0.02 | < 0.02 | < 0.02 | < 0.1 | 0.12 | 0.12 |  |  |  |  | 0.027797
| 1 | 5 | < 0.3 | < 0.5 | 0.04 | 0.04 | < 0.02 | 0.07 | < 0.05 | 0.09 | 0.04 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03449 | < 0.02 | < 0.02 | < 0.02 | < 0.1 | 0.13 | 0.12 |  |  |  |  | 31
| 1 | 6 | < 0.3 | < 0.5 | 0.05 | 0.04 | < 0.02 | 0.07 | < 0.05 | < 0.02 | 0.03 | 0.02 | 0.02 | < 0.02 | 0.02 | 0.02 | < 0.02 | 0.03326 | < 0.02 | < 0.02 | < 0.02 | < 0.1 | 0.12 | 0.14
| Mean |  |  |  | 0.048333 | 0.041667 | 0.03 | 0.071667 |  | 0.055 | 0.036667 | 0.024 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.030688 | 0.03 |  |  |  | 0.12333 | 0.13167
| Median |  |  |  | 0.05 | 0.04 | 0.03 | 0.07 |  | 0.055 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03056 | 0.03 |  |  |  | 0.12 | 0.135
| Std Dev. |  |  |  | 0.0040825 | 0.0040825 | 0 | 0.0040825 |  | 0.049497 | 0.005164 | 0.0054772 | 0 | 0 | 0 | 0 | 0 | 0.0030473 | 0 |  |  |  | 0.005164 | 0.0098319
| Rel.Std.Dev. |  |  |  | 8.447% | 9.798% | 0% | 5.696% |  | 90% | 14.08% | 22.82% | 0% | 0% | 0% | 0% | 0% | 9.93% | 0% |  |  |  | 4.187% | 7.467%
| PDM3 |  |  |  | 73.88% | 49.9% | 7.927% | 157.8% |  | 97.87% | 31.91% | -13.66% | -28.05% | -28.05% | -28.05% | -28.05% | -28.05% | 10.4% | 7.927% |  |  |  | 343.7% | 373.7%
| Z-Score (Absolute) |  | 7.22 | 13 | 1.35 | 0.96 | 0.29 | 2.7 | 0 | 0 | 0.67 | 0.19 | 0.39 | 0.48 | 0.48 | 0.48 | 0.67 | 0.33 | 0.48 | 0.87 | 0.87 | 1.44 | 5.68 | 6.17
| Analytical results for Ce in OREAS 471 (Certified Value 206 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 192 | 197.3082 | 193.2 | 202 | 162.92 | 211 | 193 | 215.35 | 211.26 | 205.62 | 218 | 208 | 218 | 211 | 205 | 196.826 | 211.16 | 203.7 | 214.45 | 205.29 | 232.7993 | 233.8233 | 217.1373 | 211.43 | 185.44 |  | 1
| 1 | 2 | 190 | 197.9286 | 194.8 | 204 | 179.99 | 210 | 191 | 200.71 | 207.45 | 200.03 | 202 | 219 | 210 | 219 | 209 | 197.6565 | 209.16 | 193.53 | 218.23 | 200.377 | 247.2406 | 233.9242 | 220.1911 | 212.96 | 186.09 |  | #N/A
| 1 | 3 | 193 | 199.8898 | 191.9 | 206 | 189.67 | 209 | 187 | 212.3 | 213.86 | 201.85 | 206 | 224 | 201 | 223 | 203 | 202.912 | 206.85 | 163.76 | 215.07 | 184.837 | 248.4367 | 234.1337 | 220.0273 | 211.18 | 187.57 |  | 16
| 1 | 4 | 196 | 198.2257 | 193.2 | 207 | 162.81 | 209 | 188 | 213.79 | 215.72 | 202.36 | 215 | 220 | 216 | 212 | 203 | 197.1766 | 211.27 | 185.26 | 220.41 | 183.219 | 226.9463 | 234.5605 | 215.409 | 213.11 | 185.58 |  | 205.9242
| 1 | 5 | 193 | 201.6402 | 192.2 | 205 | 186.13 | 209 | 187 | 220.21 | 224.53 | 205.55 | 212 | 214 | 215 | 218 | 179 | 196.2433 | 211.35 | 188.6 | 211.07 | 186 | 256.3582 | 234.4053 | 218.6645 | 211.07 | 183.75 |  | 100
| 1 | 6 | 198 | 199.8009 | 195.8 | 209 | 175.85 | 207 | 190 | 216.37 | 224.62 | 206.44 | 216 | 217 | 207 | 213 | 201 | 194.1167 | 211.76 | 180.18 | 214.61 | 178.138 | 224.0943 | 234.2967 | 217.3794 | 211.72 | 184.87
| Mean |  | 193.6667 | 199.1322 | 193.5167 | 205.5 | 176.2283 | 209.1667 | 189.3333 | 213.1217 | 216.24 | 203.6417 | 211.5 | 217 | 211.1667 | 216 | 200 | 197.4885 | 210.2583 | 185.8383 | 215.64 | 189.6435 | 239.3126 | 234.1906 | 218.1348 | 211.9117 | 185.55
| Median |  | 193 | 199.0133 | 193.2 | 205.5 | 177.92 | 209 | 189 | 214.57 | 214.79 | 203.955 | 213.5 | 218 | 212.5 | 215.5 | 203 | 197.0013 | 211.215 | 186.93 | 214.84 | 185.4185 | 240.0199 | 234.2152 | 218.022 | 211.575 | 185.51
| Std Dev. |  | 2.8752 | 1.6068 | 1.5105 | 2.429 | 11.4033 | 1.3292 | 2.4221 | 6.6473 | 7.0278 | 2.58 | 6.253 | 5.5136 | 6.4317 | 4.7329 | 10.6395 | 2.9291 | 1.903 | 13.4583 | 3.2603 | 10.6775 | 13.1423 | 0.28401 | 1.8486 | 0.89972 | 1.2718
| Rel.Std.Dev. |  | 1.485% | 0.8069% | 0.7806% | 1.182% | 6.471% | 0.6355% | 1.279% | 3.119% | 3.25% | 1.267% | 2.957% | 2.541% | 3.046% | 2.191% | 5.32% | 1.483% | 0.9051% | 7.242% | 1.512% | 5.63% | 5.492% | 0.1213% | 0.8475% | 0.4246% | 0.6854%
| PDM3 |  | -5.952% | -3.298% | -6.025% | -0.206% | -14.42% | 1.575% | -8.057% | 3.495% | 5.01% | -1.108% | 2.708% | 5.379% | 2.546% | 4.893% | -2.877% | -4.097% | 2.105% | -9.754% | 4.718% | -7.906% | 16.21% | 13.73% | 5.93% | 2.908% | -9.894%
| Z-Score (Absolute) |  | 1.17 | 0.75 | 1.18 | 0.28 | 2.48 | 0 | 1.49 | 0.3 | 0.53 | 0.42 | 0.18 | 0.59 | 0.15 | 0.51 | 0.69 | 0.88 | 0.08 | 1.75 | 0.49 | 1.47 | 2.27 | 1.88 | 0.67 | 0.21 | 1.78
| Analytical results for Co in OREAS 471 (Certified Value 15.5 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 16 | 14.9188 | 18.3 | 14.7 | 14.5 | 14.3 | 16 | 16.2 | 16 | 15.9 | 15.2 | 16.9 | 16.4 | 14.2 | 14.4 | 15.4501 | 15.7 | 15.4 | 15.4 | 14.721 | 13.6428 | 15.64 | 16.9631 | 15.3 |  |  | 1
| 1 | 2 | 15 | 14.9917 | 18.4 | 14.7 | 15.7 | 14.8 | 16 | 16.2 | 16 | 15.8 | 15.6 | 17.4 | 15.6 | 15.4 | 15.4 | 14.9471 | 15.9 | 15.8 | 15.7 | 16.107 | 13.8069 | 15.31 | 17.3109 | 15.5 |  |  | #N/A
| 1 | 3 | 15 | 15.2283 | 17.9 | 14.3 | 14.8 | 14.4 | 16 | 15.7 | 15.7 | 17 | 15.2 | 17.2 | 15.4 | 15.3 | 14.7 | 15.7096 | 15.6 | 16.1 | 15.4 | 16.264 | 14.0015 | 15.41 | 16.7504 | 15.7 |  |  | 16
| 1 | 4 | 16 | 14.9876 | 17.8 | 14.5 | 14.4 | 14.6 | 15 | 15.8 | 17.9 | 15.7 | 15.3 | 16.9 | 15.9 | 14.4 | 15.2 | 14.6887 | 16.3 | 16.1 | 15.4 | 15.472 | 13.662 | 16.11 | 16.9561 | 15.5 |  |  | 15.5266
| 1 | 5 | 14 | 15.1657 | 17.9 | 14.7 | 14.9 | 14.6 | 16 | 16 | 16.5 | 15.5 | 15.8 | 16.6 | 15.8 | 14.8 | 14.8 | 14.8726 | 15.7 | 15.3 | 15.5 | 15.752 | 13.8452 | 15.52 | 17.2265 | 15.3 |  |  | 101
| 1 | 6 | 16 | 15.5157 | 18.6 | 13.8 | 14.8 | 14.3 | 16 | 15.9 | 16.4 | 15.7 | 16.1 | 16.6 | 15.7 | 14.6 | 15.6 | 14.7127 | 15.8 | 15.1 | 15.7 | 14.749 | 13.9106 | 15.48 | 17.9708 | 15.2
| Mean |  | 15.3333 | 15.1346 | 18.15 | 14.45 | 14.85 | 14.5 | 15.8333 | 15.9667 | 16.4167 | 15.9333 | 15.5333 | 16.9333 | 15.8 | 14.7833 | 15.0167 | 15.0635 | 15.8333 | 15.6333 | 15.5167 | 15.5108 | 13.8115 | 15.5783 | 17.1963 | 15.4167
| Median |  | 15.5 | 15.0787 | 18.1 | 14.6 | 14.8 | 14.5 | 16 | 15.95 | 16.2 | 15.75 | 15.45 | 16.9 | 15.75 | 14.7 | 15 | 14.9098 | 15.75 | 15.6 | 15.45 | 15.612 | 13.826 | 15.5 | 17.0948 | 15.4
| Std Dev. |  | 0.8165 | 0.2207 | 0.32711 | 0.35637 | 0.45935 | 0.2 | 0.40825 | 0.20656 | 0.78337 | 0.53914 | 0.36697 | 0.32042 | 0.34059 | 0.48339 | 0.45789 | 0.4196 | 0.25033 | 0.4274 | 0.1472 | 0.6613 | 0.13992 | 0.2828 | 0.43006 | 0.18348
| Rel.Std.Dev. |  | 5.325% | 1.458% | 1.802% | 2.466% | 3.093% | 1.379% | 2.578% | 1.294% | 4.772% | 3.384% | 2.362% | 1.892% | 2.156% | 3.27% | 3.049% | 2.786% | 1.581% | 2.734% | 0.9486% | 4.263% | 1.013% | 1.815% | 2.501% | 1.19%
| PDM3 |  | -1.245% | -2.525% | 16.9% | -6.934% | -4.358% | -6.612% | 1.976% | 2.834% | 5.733% | 2.62% | 0.04337% | 9.06% | 1.761% | -4.787% | -3.284% | -2.983% | 1.976% | 0.6874% | -0.06397% | -0.1015% | -11.05% | 0.3332% | 10.75% | -0.708%
| Z-Score (Absolute) |  | NA | 0.58 | 3.92 | 1.61 | 1.01 | 1.53 | NA | 0.66 | 1.33 | 0.61 | 0.01 | 2.1 | 0.41 | 1.11 | 0.76 | 0.69 | 0.46 | 0.16 | 0.01 | 0.02 | 2.56 | 0.08 | 2.5 | 0.16
Results from laboratories 1 and 8 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 471 (Certified Value 51 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
1
| 1 | 1 | 43 | 49.8665 | 53.9 | 43.6 | 23 | 52 | 52 | 55 | 45 | 56 | 53 | 55 | 52 | 51 | 52 | 57.395 | 56 | 43 | 45 | 49.848 | 46.3255 | 51 | 43 |  |  |  | 1
| 1 | 2 | 43 | 49.5753 | 54.7 | 45 | 22 | 51 | 52 | 52 | 56 | 53 | 53 | 53 | 50 | 51 | 52 | 58.2565 | 56 | 45 | 46 | 45.169 | 46.4485 | 50 | 50 |  |  |  | #N/A
| 1 | 3 | 43 | 50.4229 | 53.5 | 45.7 | 20 | 54 | 53 | 53 | 54 | 53 | 50 | 56 | 51 | 50 | 50 | 58.2214 | 55 | 46 | 44 | 76.424 | 46.187 | 50 | 44 |  |  |  | 16
| 1 | 4 | 44 | 50.9555 | 54.2 | 49.5 | 24 | 52 | 53 | 53 | 50 | 55 | 54 | 59 | 51 | 51 | 51 | 57.0645 | 56 | 46 | 44 | 68.399 | 46.839 | 51 | 47 |  |  |  | 50.5618
| 1 | 5 | 44 | 49.4316 | 53.6 | 46.6 | 25 | 50 | 51 | 54 | 48 | 55 | 51 | 56 | 50 | 50 | 49 | 56.0908 | 54 | 42 | 42 | 52.791 | 46.8875 | 50 | 48 |  |  |  | 102
| 1 | 6 | 40 | 50.1298 | 54.4 | 46.1 | 26 | 54 | 51 | 51 | 43 | 60 | 54 | 54 | 51 | 52 | 52 | 56.7315 | 55 | 46 | 46 | 58.491 | 46.68 | 50 | 44
| Mean |  | 42.8333 | 50.0636 | 54.05 | 46.0833 | 23.3333 | 52.1667 | 52 | 53 | 49.3333 | 55.3333 | 52.5 | 55.5 | 50.8333 | 50.8333 | 51 | 57.2933 | 55.3333 | 44.6667 | 44.5 | 58.5203 | 46.5613 | 50.3333 | 46
| Median |  | 43 | 49.9982 | 54.05 | 45.9 | 23.5 | 52 | 52 | 53 | 49 | 55 | 53 | 55.5 | 51 | 51 | 51.5 | 57.2298 | 55.5 | 45.5 | 44.5 | 55.641 | 46.5643 | 50 | 45.5
| Std Dev. |  | 1.472 | 0.56668 | 0.46797 | 1.9712 | 2.1602 | 1.6021 | 0.89443 | 1.4142 | 5.0465 | 2.582 | 1.6432 | 2.0736 | 0.75277 | 0.75277 | 1.2649 | 0.85009 | 0.8165 | 1.7512 | 1.5166 | 11.8695 | 0.28497 | 0.5164 | 2.7568
| Rel.Std.Dev. |  | 3.436% | 1.132% | 0.8658% | 4.277% | 9.258% | 3.071% | 1.72% | 2.668% | 10.23% | 4.666% | 3.13% | 3.736% | 1.481% | 1.481% | 2.48% | 1.484% | 1.476% | 3.921% | 3.408% | 20.28% | 0.612% | 1.026% | 5.993%
| PDM3 |  | -15.29% | -0.9853% | 6.899% | -8.857% | -53.85% | 3.174% | 2.844% | 4.822% | -2.43% | 9.437% | 3.833% | 9.767% | 0.537% | 0.537% | 0.8667% | 13.31% | 9.437% | -11.66% | -11.99% | 15.74% | -7.912% | -0.4519% | -9.022%
| Z-Score (Absolute) |  | 1.26 | 0.12 | 0.51 | 0.75 | 4.34 | 0.21 | 0.18 | 0.34 | 0.24 | 0.71 | 0.26 | 0.74 | 0 | 0 | 0.03 | 1.02 | 0.71 | 0.97 | 1 | 1.21 | 0.67 | 0.08 | 0.76
| Analytical results for Cs in OREAS 471 (Certified Value 6.99 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 6.59 | 6.7172 | 7.3 | 6.98 | 6 | 6.7 | 6.9 | 7.05 | 7.06 | 7.16 | 6.9 | 7.32 | 7.4 | 6.91 | 6.9 | 7.1016 | 7.06 | 6 | 7.4 | 7.18 | 6.3226 | 7.2 | 6.9159 | 6.71 |  |  | 1
| 1 | 2 | 6.48 | 6.7856 | 7.1 | 6.97 | 7 | 6.99 | 6.7 | 6.36 | 7.48 | 7.04 | 7.01 | 7.37 | 7.08 | 7.19 | 7.19 | 6.9887 | 7.01 | 6 | 7.5 | 7.057 | 6.128 | 7.17 | 7.0481 | 6.54 |  |  | #N/A
| 1 | 3 | 6.62 | 6.5545 | 6.7 | 6.81 | 7 | 6.86 | 6.7 | 6.81 | 7.59 | 7.15 | 6.84 | 7.64 | 6.93 | 7.26 | 6.98 | 7.0863 | 6.98 | 5.6 | 7.3 | 7.067 | 6.2573 | 7.22 | 6.915 | 6.73 |  |  | 16
| 1 | 4 | 6.62 | 6.6269 | 7.1 | 6.98 | 6 | 6.95 | 6.8 | 6.9 | 7.56 | 7.19 | 7.09 | 7.6 | 7.27 | 6.98 | 6.89 | 7.0522 | 7.09 | 6 | 7.5 | 7.292 | 6.4433 | 7.19 | 6.6451 | 6.53 |  |  | 6.9862
| 1 | 5 | 6.69 | 6.6321 | 6.6 | 6.93 | 6 | 6.98 | 6.6 | 7.01 | 7.73 | 7.1 | 6.98 | 7.53 | 7.22 | 7.14 | 6.54 | 6.9214 | 7.09 | 5.8 | 7.3 | 7.182 | 6.2186 | 7.22 | 6.8564 | 6.49 |  |  | 103
| 1 | 6 | 6.78 | 6.8197 | 6.8 | 6.59 | 6 | 6.79 | 6.8 | 7.02 | 7.9 | 7.26 | 7.08 | 7.71 | 7.23 | 6.99 | 6.88 | 6.8212 | 7.07 | 5.6 | 7.4 | 7.11 | 6.226 | 7.18 | 6.8645 | 6.69
| Mean |  | 6.63 | 6.6893 | 6.9333 | 6.8767 | 6.3333 | 6.8783 | 6.75 | 6.8583 | 7.5533 | 7.15 | 6.9833 | 7.5283 | 7.1883 | 7.0783 | 6.8967 | 6.9952 | 7.05 | 5.8333 | 7.4 | 7.148 | 6.266 | 7.1967 | 6.8742 | 6.615
| Median |  | 6.62 | 6.6746 | 6.95 | 6.95 | 6 | 6.905 | 6.75 | 6.955 | 7.575 | 7.155 | 6.995 | 7.565 | 7.225 | 7.065 | 6.895 | 7.0204 | 7.065 | 5.9 | 7.4 | 7.145 | 6.2416 | 7.195 | 6.8898 | 6.615
| Std Dev. |  | 0.1004 | 0.10235 | 0.27325 | 0.15462 | 0.5164 | 0.11652 | 0.10488 | 0.26011 | 0.28338 | 0.075366 | 0.098928 | 0.15433 | 0.1629 | 0.1379 | 0.21002 | 0.10829 | 0.045166 | 0.19664 | 0.089443 | 0.088546 | 0.10734 | 0.020656 | 0.13156 | 0.10616
| Rel.Std.Dev. |  | 1.514% | 1.53% | 3.941% | 2.248% | 8.154% | 1.694% | 1.554% | 3.793% | 3.752% | 1.054% | 1.417% | 2.05% | 2.266% | 1.948% | 3.045% | 1.548% | 0.6407% | 3.371% | 1.209% | 1.239% | 1.713% | 0.287% | 1.914% | 1.605%
| PDM3 |  | -5.099% | -4.249% | -0.7567% | -1.568% | -9.345% | -1.544% | -3.381% | -1.83% | 8.118% | 2.345% | -0.04099% | 7.76% | 2.893% | 1.319% | -1.282% | 0.1289% | 0.9133% | -16.5% | 5.923% | 2.316% | -10.31% | 3.013% | -1.604% | -5.313%
| Z-Score (Absolute) |  | 0.95 | 0.77 | 0 | 0.18 | NA | 0.17 | 0.58 | 0.24 | 1.95 | 0.68 | 0.16 | 1.87 | 0.8 | 0.46 | 0.12 | 0.19 | 0.37 | 3.46 | 1.47 | 0.67 | 2.1 | 0.83 | 0.19 | 1
Results from laboratory 6 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 471 (Certified Value 32.7 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 31 | 31.2423 | 25.2 | 34.2 | 33.7 | 30.7 | 29 | 33.6 | 33.5 | 32.8 | 33.5 | 32.1 | 33.7 | 32.5 | 32.3 | 32.1521 | 32 | 33.3 | 31.8 | 33.35 | 31.2 | 32.1 | 29.9624 | 37.3 |  |  | 1
| 1 | 2 | 30 | 32.0138 | 23.8 | 34 | 35.4 | 34.1 | 29 | 33.1 | 32.9 | 31.7 | 33.6 | 34.3 | 32.3 | 33.5 | 33.3 | 31.6637 | 32 | 34.9 | 32.2 | 35.513 | 31.36 | 32.5 | 30.8604 | 37.1 |  |  | #N/A
| 1 | 3 | 32 | 33.2161 | 21.1 | 34.4 | 32.2 | 37.4 | 27 | 32.5 | 34.1 | 32.1 | 32.8 | 33.3 | 31.1 | 33.5 | 32.2 | 33.1045 | 32 | 33.9 | 32.1 | 35.619 | 31.62 | 32.6 | 30.8773 | 36.4 |  |  | 16
| 1 | 4 | 32 | 31.9837 | 20.5 | 34.9 | 32.4 | 32.7 | 29 | 34 | 34.8 | 32.9 | 33.2 | 33 | 32.7 | 32.6 | 32.7 | 32.2048 | 33 | 34.9 | 32 | 34.875 | 31.47 | 32.1 | 30.7482 | 37.5 |  |  | 32.7172
| 1 | 5 | 30 | 31.8829 | 20.4 | 34 | 34.4 | 33.2 | 28 | 33.7 | 33.5 | 32.5 | 33.3 | 33.2 | 32.7 | 33.4 | 31.6 | 31.5375 | 32 | 35.1 | 32 | 34.438 | 31.47 | 32.1 | 30.2041 | 36.5 |  |  | 104
| 1 | 6 | 32 | 32.5207 | 24.6 | 33.6 | 33.8 | 31.3 | 29 | 33.5 | 32.4 | 32.4 | 34.3 | 32.7 | 32.4 | 33 | 32.9 | 32.1224 | 33 | 34.2 | 32.4 | 32.614 | 31.49 | 32.6 | 30.1567 | 37.3
| Mean |  | 31.1667 | 32.1432 | 22.6 | 34.1833 | 33.65 | 33.2333 | 28.5 | 33.4 | 33.5333 | 32.4 | 33.45 | 33.1 | 32.4833 | 33.0833 | 32.5 | 32.1308 | 32.3333 | 34.3833 | 32.0833 | 34.4015 | 31.435 | 32.3333 | 30.4682 | 37.0167
| Median |  | 31.5 | 31.9988 | 22.45 | 34.1 | 33.75 | 32.95 | 29 | 33.55 | 33.5 | 32.45 | 33.4 | 33.1 | 32.55 | 33.2 | 32.5 | 32.1372 | 32 | 34.55 | 32.05 | 34.6565 | 31.47 | 32.3 | 30.4761 | 37.2
| Std Dev. |  | 0.98319 | 0.66562 | 2.1772 | 0.44008 | 1.2095 | 2.3897 | 0.83666 | 0.52915 | 0.8501 | 0.44721 | 0.501 | 0.72938 | 0.84004 | 0.4535 | 0.59666 | 0.55212 | 0.5164 | 0.70545 | 0.20412 | 1.2037 | 0.14181 | 0.2582 | 0.40552 | 0.45789
| Rel.Std.Dev. |  | 3.155% | 2.071% | 9.633% | 1.287% | 3.594% | 7.191% | 2.936% | 1.584% | 2.535% | 1.38% | 1.498% | 2.204% | 2.586% | 1.371% | 1.836% | 1.718% | 1.597% | 2.052% | 0.6362% | 3.499% | 0.4511% | 0.7986% | 1.331% | 1.237%
| PDM3 |  | -4.739% | -1.754% | -30.92% | 4.481% | 2.851% | 1.577% | -12.89% | 2.087% | 2.494% | -0.9697% | 2.24% | 1.17% | -0.7149% | 1.119% | -0.664% | -1.792% | -1.173% | 5.092% | -1.938% | 5.148% | -3.919% | -1.173% | -6.874% | 13.14%
| Z-Score (Absolute) |  | 0.96 | 0.25 | 7.15 | 1.22 | 0.84 | 0.54 | 2.88 | 0.66 | 0.75 | 0.07 | 0.69 | 0.44 | 0.01 | 0.43 | 0.01 | 0.26 | 0.11 | 1.37 | 0.3 | 1.38 | 0.76 | 0.11 | 1.46 | 3.27
| Analytical results for Dy in OREAS 471 (Certified Value 4.85 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 4.9 | 4.4413 | 4.81 | 4.65 | 5.0508 | 4.99 | 4.6 | 4.93 | 4.94 | 5.07 | 4.99 | 4.99 | 4.44 | 4.44 | 4.48 | 4.8426 | 4.86 | 4.8 | 5 | 4.621 | 5.2303 | 4.5356 | 4.8767 | 5.04 | 5 |  | 1
| 1 | 2 | 4.8 | 4.6098 | 4.68 | 4.61 | 5.5035 | 5.22 | 4.6 | 4.63 | 5.14 | 4.99 | 4.77 | 5.12 | 4.32 | 4.62 | 4.72 | 4.955 | 4.92 | 4.5 | 5.1 | 4.933 | 4.825 | 4.4546 | 4.9807 | 5.13 | 5.05 |  | #N/A
| 1 | 3 | 4.5 | 4.6103 | 4.39 | 4.7 | 5.4661 | 5.1 | 4.7 | 4.88 | 5.31 | 5.03 | 4.93 | 5.16 | 4.19 | 4.69 | 4.5 | 5.1324 | 4.81 | 4.3 | 5 | 5.179 | 5.5587 | 4.5472 | 5.0379 | 5.14 | 4.93 |  | 16
| 1 | 4 | 4.5 | 4.521 | 4.69 | 4.69 | 4.8597 | 4.96 | 4.6 | 4.84 | 5.38 | 5.4 | 4.97 | 5.24 | 4.4 | 4.48 | 4.43 | 4.9815 | 4.89 | 4.5 | 5.1 | 5.373 | 5.3345 | 4.5558 | 4.9046 | 4.96 | 4.93 |  | 4.8526
| 1 | 5 | 4.7 | 4.6018 | 4.29 | 4.81 | 5.4284 | 4.99 | 4.7 | 5.76 | 5.58 | 5.21 | 4.84 | 5.1 | 4.35 | 4.59 | 4.29 | 4.9434 | 4.98 | 4.5 | 5 | 5.487 | 5.6906 | 4.5715 | 4.8879 | 5.01 | 4.97 |  | 105
| 1 | 6 | 4.7 | 4.7115 | 4.5 | 4.63 | 5.1425 | 5.08 | 4.6 | 4.96 | 5.58 | 5.03 | 4.98 | 5.08 | 4.33 | 4.47 | 4.49 | 5.0467 | 4.94 | 4.2 | 5 | 4.627 | 4.9967 | 4.5228 | 4.9484 | 5.07 | 4.9
| Mean |  | 4.6833 | 4.5826 | 4.56 | 4.6817 | 5.2418 | 5.0567 | 4.6333 | 5 | 5.3217 | 5.1217 | 4.9133 | 5.115 | 4.3383 | 4.5483 | 4.485 | 4.9836 | 4.9 | 4.4667 | 5.0333 | 5.0367 | 5.2726 | 4.5312 | 4.9394 | 5.0583 | 4.9633
| Median |  | 4.7 | 4.6058 | 4.59 | 4.67 | 5.2854 | 5.035 | 4.6 | 4.905 | 5.345 | 5.05 | 4.95 | 5.11 | 4.34 | 4.535 | 4.485 | 4.9683 | 4.905 | 4.5 | 5 | 5.056 | 5.2824 | 4.5414 | 4.9265 | 5.055 | 4.95
| Std Dev. |  | 0.16021 | 0.091931 | 0.1996 | 0.071671 | 0.26304 | 0.097297 | 0.05164 | 0.39013 | 0.25127 | 0.15626 | 0.089144 | 0.083367 | 0.085654 | 0.099482 | 0.13896 | 0.098437 | 0.060332 | 0.20656 | 0.05164 | 0.37082 | 0.32822 | 0.041113 | 0.062104 | 0.06969 | 0.055015
| Rel.Std.Dev. |  | 3.421% | 2.006% | 4.377% | 1.531% | 5.018% | 1.924% | 1.115% | 7.803% | 4.722% | 3.051% | 1.814% | 1.63% | 1.974% | 2.187% | 3.098% | 1.975% | 1.231% | 4.624% | 1.026% | 7.362% | 6.225% | 0.9073% | 1.257% | 1.378% | 1.108%
| PDM3 |  | -3.489% | -5.565% | -6.03% | -3.523% | 8.02% | 4.205% | -4.519% | 3.037% | 9.665% | 5.544% | 1.251% | 5.407% | -10.6% | -6.271% | -7.576% | 2.699% | 0.976% | -7.954% | 3.724% | 3.792% | 8.655% | -6.623% | 1.787% | 4.239% | 2.281%
| Z-Score (Absolute) |  | 0.67 | 0.94 | 1 | 0.68 | 0.8 | 0.31 | 0.81 | 0.16 | 1.01 | 0.48 | 0.07 | 0.46 | 1.58 | 1.03 | 1.2 | 0.12 | 0.1 | 1.24 | 0.25 | 0.26 | 0.88 | 1.07 | 0 | 0.31 | 0.06
| Analytical results for Er in OREAS 471 (Certified Value 2.3 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 2.1 | 2.0873 | 2.45 | 2.3 | 2.3779 | 2.2 | 2.3 | 2.34 | 2.2 | 2.41 | 2.24 | 2.23 | 2.11 | 2 | 2.16 | 2.4198 | 2.84 | 2.2 | 2.3 | 2.297 | 2.2731 | 2.521 | 2.2779 | 2.39 | 2.41 |  | 1
| 1 | 2 | 2.2 | 2.0785 | 2.35 | 2.25 | 2.5981 | 2.18 | 2.3 | 2.24 | 2.32 | 2.4 | 2.19 | 2.32 | 2.04 | 2.11 | 2.25 | 2.5008 | 2.33 | 2.2 | 2.3 | 2.455 | 2.0968 | 2.5755 | 2.3155 | 2.4 | 2.45 |  | #N/A
| 1 | 3 | 2.2 | 2.1637 | 2.18 | 2.31 | 2.5892 | 2.26 | 2.4 | 2.22 | 2.33 | 2.43 | 2.23 | 2.34 | 2 | 2.15 | 2.13 | 2.4933 | 2.3 | 2 | 2.3 | 2.531 | 2.7618 | 2.5384 | 2.3264 | 2.43 | 2.43 |  | 16
| 1 | 4 | 2 | 2.1833 | 2.37 | 2.3 | 2.3241 | 2.2 | 2.4 | 2.27 | 2.41 | 2.39 | 2.22 | 2.4 | 2.08 | 2.08 | 2.08 | 2.4657 | 2.36 | 2.1 | 2.4 | 2.676 | 2.3025 | 2.4807 | 2.2196 | 2.42 | 2.39 |  | 2.2995
| 1 | 5 | 2.3 | 2.1551 | 2.19 | 2.36 | 2.5604 | 2.2 | 2.3 | 2.34 | 2.5 | 2.41 | 2.17 | 2.28 | 2.06 | 2.11 | 2.05 | 2.3868 | 2.37 | 2.2 | 2.3 | 2.695 | 2.8409 | 2.5017 | 2.3008 | 2.41 | 2.43 |  | 106
| 1 | 6 | 2.1 | 2.1939 | 2.26 | 2.27 | 2.4493 | 2.17 | 2.3 | 2.3 | 2.57 | 2.48 | 2.27 | 2.28 | 2.04 | 2.09 | 2.11 | 2.3488 | 2.39 | 2 | 2.3 | 2.243 | 2.2164 | 2.473 | 2.2922 | 2.46 | 2.45
| Mean |  | 2.15 | 2.1436 | 2.3 | 2.2983 | 2.4832 | 2.2017 | 2.3333 | 2.285 | 2.3883 | 2.42 | 2.22 | 2.3083 | 2.055 | 2.09 | 2.13 | 2.4359 | 2.4317 | 2.1167 | 2.3167 | 2.4828 | 2.4153 | 2.515 | 2.2887 | 2.4183 | 2.4267
| Median |  | 2.15 | 2.1594 | 2.305 | 2.3 | 2.5048 | 2.2 | 2.3 | 2.285 | 2.37 | 2.41 | 2.225 | 2.3 | 2.05 | 2.1 | 2.12 | 2.4427 | 2.365 | 2.15 | 2.3 | 2.493 | 2.2878 | 2.5113 | 2.2965 | 2.415 | 2.43
| Std Dev. |  | 0.10488 | 0.04911 | 0.1077 | 0.037639 | 0.1166 | 0.031252 | 0.05164 | 0.050498 | 0.13378 | 0.032249 | 0.035777 | 0.058793 | 0.037815 | 0.0502 | 0.070143 | 0.061077 | 0.20253 | 0.098319 | 0.040825 | 0.1884 | 0.30828 | 0.038357 | 0.03794 | 0.024833 | 0.023381
| Rel.Std.Dev. |  | 4.878% | 2.291% | 4.683% | 1.638% | 4.696% | 1.419% | 2.213% | 2.21% | 5.601% | 1.333% | 1.612% | 2.547% | 1.84% | 2.402% | 3.293% | 2.507% | 8.329% | 4.645% | 1.762% | 7.588% | 12.76% | 1.525% | 1.658% | 1.027% | 0.9635%
| PDM3 |  | -6.502% | -6.779% | 0.0214% | -0.05108% | 7.986% | -4.255% | 1.471% | -0.6309% | 3.863% | 5.24% | -3.458% | 0.3838% | -10.63% | -9.111% | -7.371% | 5.93% | 5.747% | -7.951% | 0.7462% | 7.972% | 5.034% | 9.373% | -0.4686% | 5.167% | 5.53%
| Z-Score (Absolute) |  | 0.96 | 0.99 | 0.05 | 0.06 | 1.06 | 0.64 | 0.15 | 0.14 | 0.48 | 0.67 | 0.53 | 0 | 1.53 | 1.32 | 1.08 | 0.77 | 0.74 | 1.16 | 0.05 | 1.05 | 0.65 | 1.25 | 0.12 | 0.66 | 0.71
| Analytical results for Eu in OREAS 471 (Certified Value 2.63 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 2.41 | 2.4607 | 3.14 | 2.61 | 2.627 | 2.73 | 2.4 | 2.66 | 2.77 | 2.65 | 2.66 | 2.71 | 2.57 | 2.44 | 2.39 | 2.4718 | 2.69 | 2.7 | 2.5 | 2.661 | 3.2618 | 3.2954 | 2.6398 | 2.75 | 2.56 |  | 1
| 1 | 2 | 2.48 | 2.5586 | 3.03 | 2.54 | 2.9021 | 2.69 | 2.5 | 2.45 | 2.86 | 2.69 | 2.57 | 2.86 | 2.43 | 2.51 | 2.45 | 2.5438 | 2.7 | 2.5 | 2.6 | 2.709 | 3.2157 | 3.2848 | 2.6903 | 2.67 | 2.6 |  | #N/A
| 1 | 3 | 2.65 | 2.4974 | 2.81 | 2.63 | 2.9154 | 2.69 | 2.4 | 2.61 | 3.01 | 2.6 | 2.57 | 2.86 | 2.41 | 2.55 | 2.41 | 2.5078 | 2.71 | 2.3 | 2.5 | 2.863 | 3.6375 | 3.2882 | 2.7363 | 2.78 | 2.51 |  | 16
| 1 | 4 | 2.53 | 2.4613 | 3.07 | 2.59 | 2.5993 | 2.73 | 2.5 | 2.61 | 2.98 | 2.63 | 2.63 | 2.9 | 2.47 | 2.44 | 2.34 | 2.4158 | 2.7 | 2.5 | 2.6 | 2.815 | 3.3668 | 3.3189 | 2.5363 | 2.64 | 2.53 |  | 2.6298
| 1 | 5 | 2.5 | 2.5443 | 2.78 | 2.63 | 2.8822 | 2.69 | 2.5 | 2.71 | 3.1 | 2.7 | 2.57 | 2.79 | 2.49 | 2.46 | 2.21 | 2.4101 | 2.75 | 2.5 | 2.5 | 2.865 | 3.6647 | 3.2926 | 2.6327 | 2.67 | 2.58 |  | 107
| 1 | 6 | 2.64 | 2.5465 | 2.88 | 2.61 | 2.7288 | 2.76 | 2.5 | 2.66 | 3.14 | 2.69 | 2.69 | 2.83 | 2.45 | 2.42 | 2.37 | 2.4533 | 2.94 | 2.4 | 2.5 | 2.634 | 3.771 | 3.3149 | 2.6911 | 2.74 | 2.55
| Mean |  | 2.535 | 2.5115 | 2.9517 | 2.6017 | 2.7758 | 2.715 | 2.4667 | 2.6167 | 2.9767 | 2.66 | 2.615 | 2.825 | 2.47 | 2.47 | 2.3617 | 2.4671 | 2.7483 | 2.4833 | 2.5333 | 2.7578 | 3.4862 | 3.2991 | 2.6544 | 2.7083 | 2.555
| Median |  | 2.515 | 2.5209 | 2.955 | 2.61 | 2.8055 | 2.71 | 2.5 | 2.635 | 2.995 | 2.67 | 2.6 | 2.845 | 2.46 | 2.45 | 2.38 | 2.4626 | 2.705 | 2.5 | 2.5 | 2.762 | 3.5021 | 3.294 | 2.665 | 2.705 | 2.555
| Std Dev. |  | 0.093968 | 0.044329 | 0.14851 | 0.033714 | 0.14302 | 0.029496 | 0.05164 | 0.089815 | 0.14095 | 0.04 | 0.05282 | 0.067157 | 0.056569 | 0.0498 | 0.083046 | 0.052195 | 0.096212 | 0.13292 | 0.05164 | 0.10287 | 0.23392 | 0.014289 | 0.069249 | 0.055648 | 0.032711
| Rel.Std.Dev. |  | 3.707% | 1.765% | 5.032% | 1.296% | 5.152% | 1.086% | 2.094% | 3.432% | 4.735% | 1.504% | 2.02% | 2.377% | 2.29% | 2.016% | 3.516% | 2.116% | 3.501% | 5.352% | 2.038% | 3.73% | 6.71% | 0.4331% | 2.609% | 2.055% | 1.28%
| PDM3 |  | -3.605% | -4.498% | 12.24% | -1.07% | 5.553% | 3.24% | -6.203% | -0.4993% | 13.19% | 1.149% | -0.5627% | 7.423% | -6.076% | -6.076% | -10.2% | -6.187% | 4.507% | -5.569% | -3.668% | 4.869% | 32.57% | 25.45% | 0.9359% | 2.986% | -2.844%
| Z-Score (Absolute) |  | 0.41 | 0.53 | 1.69 | 0.08 | 0.8 | 0.5 | 0.76 | 0 | 1.82 | 0.22 | 0.01 | 1.05 | 0.74 | 0.74 | 1.29 | 0.76 | 0.67 | 0.67 | 0.42 | 0.71 | 4.4 | 3.45 | 0.19 | 0.46 | 0.31
| Analytical results for Fe in OREAS 471 (Certified Value 4.23 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 4.26 | 4.0611 | 4.1788 | 4.49 | 3.98 | 4.11 | 4.25 | 4.24 | 3.96 | 4.3 | 4.38 | 4.08 | 4.33 | 4.11 | 4.34 | 4.3093 | 4.42 | 4.31 | 3.9429 | 4.4842 | 4.37 | 0.0029962 | 4.36 |  |  |  | 1
| 1 | 2 | 4.14 | 4.1192 | 4.1807 | 4.59 | 4.37 | 4.27 | 4.19 | 4.37 | 4.01 | 4.19 | 4.3 | 4.17 | 4.06 | 4.29 | 4.42 | 4.3326 | 4.32 | 4.34 | 3.9845 | 4.493 | 4.36 | 0.003086 | 4.4 |  |  |  | #N/A
| 1 | 3 | 4.23 | 4.1126 | 4.1609 | 4.61 | 4.02 | 4.22 | 4.2 | 4.14 | 4.04 | 4.25 | 4.17 | 4.24 | 4.18 | 4.29 | 4.19 | 4.2585 | 4.25 | 4.3 | 3.9728 | 4.4534 | 4.34 | 0.0030877 | 4.33 |  |  |  | 16
| 1 | 4 | 4.33 | 4.0907 | 4.2253 | 4.63 | 3.96 | 4.29 | 4.26 | 4.3 | 4.15 | 4.24 | 4.36 | 4.43 | 4.25 | 4.1 | 4.23 | 4.3013 | 4.32 | 4.36 | 4.03 | 4.5083 | 4.34 | 0.0030748 | 4.39 |  |  |  | 4.2336
| 1 | 5 | 4.17 | 4.15 | 4.1667 | 4.55 | 4.14 | 4.08 | 4.25 | 4.22 | 4.12 | 4.22 | 4.24 | 4.19 | 4.18 | 4.19 | 4.08 | 4.2784 | 4.23 | 4.34 | 4.0469 | 4.5014 | 4.32 | 0.0030204 | 4.42 |  |  |  | 108
| 1 | 6 | 4.23 | 4.2069 | 4.2176 | 4.59 | 4.14 | 4.14 | 4.21 | 4.29 | 3.94 | 4.28 | 4.42 | 4.14 | 4.29 | 4.21 | 4.28 | 4.3458 | 4.25 | 4.3 | 3.8363 | 4.467 | 4.36 | 0.0030157 | 4.36
| Mean |  | 4.2267 | 4.1234 | 4.1883 | 4.5767 | 4.1017 | 4.185 | 4.2267 | 4.26 | 4.0367 | 4.2467 | 4.3117 | 4.2083 | 4.215 | 4.1983 | 4.2567 | 4.3043 | 4.2983 | 4.325 | 3.9689 | 4.4845 | 4.3483 | 0.0030468 | 4.3767
| Median |  | 4.23 | 4.1159 | 4.1798 | 4.59 | 4.08 | 4.18 | 4.23 | 4.265 | 4.025 | 4.245 | 4.33 | 4.18 | 4.215 | 4.2 | 4.255 | 4.3053 | 4.285 | 4.325 | 3.9787 | 4.4886 | 4.35 | 0.0030476 | 4.375
| Std Dev. |  | 0.067132 | 0.050531 | 0.026806 | 0.050067 | 0.15263 | 0.08735 | 0.030111 | 0.07874 | 0.084538 | 0.039833 | 0.093897 | 0.1209 | 0.096488 | 0.083046 | 0.11877 | 0.032636 | 0.070828 | 0.0251 | 0.075279 | 0.020967 | 0.018348 | 4.0552e-05 | 0.03266
| Rel.Std.Dev. |  | 1.588% | 1.225% | 0.64% | 1.094% | 3.721% | 2.087% | 0.7124% | 1.848% | 2.094% | 0.938% | 2.178% | 2.873% | 2.289% | 1.978% | 2.79% | 0.7582% | 1.648% | 0.5803% | 1.897% | 0.4675% | 0.422% | 1.331% | 0.7462%
| PDM3 |  | -0.163% | -2.601% | -1.068% | 8.104% | -3.116% | -1.147% | -0.163% | 0.6244% | -4.651% | 0.3094% | 1.845% | -0.596% | -0.4385% | -0.8322% | 0.5457% | 1.671% | 1.53% | 2.16% | -6.252% | 5.928% | 2.711% | -99.93% | 3.38%
| Z-Score (Absolute) |  | 0.1 | 1.18 | 0.5 | 3.55 | 1.41 | 0.54 | 0.1 | 0.24 | 2.09 | 0.1 | 0.78 | 0.3 | 0.23 | 0.4 | 0.21 | 0.71 | 0.64 | 0.92 | 2.79 | 2.58 | 1.16 | NA | 1.46
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Ga in OREAS 471 (Certified Value 21.3 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 24.1 | 20.8399 | 21.83 | 26.2 | 20.1 | 22.6 | 22 | 21.22 | 22.32 | 21.93 | 21.3 | 21.4 | 22.1 | 20.5 | 20.7 | 21.1842 | 20.97 | 20.72 | 20.48 | 20.307 | 21.721 | 22.11 | 25.9223 | 20.16 | 18.82 |  | 1
| 1 | 2 | 24 | 21.7091 | 21.93 | 26.6 | 21.5 | 22.9 | 20.9 | 20.58 | 22.62 | 21.46 | 21.5 | 23.1 | 20.9 | 21.5 | 21.6 | 21.2087 | 20.87 | 21.52 | 20.7 | 20.679 | 21.922 | 22.48 | 25.9196 | 19.97 | 19.35 |  | #N/A
| 1 | 3 | 23.7 | 21.9216 | 21.55 | 25.4 | 20.2 | 22.2 | 21.1 | 20.28 | 22.03 | 21.76 | 20.6 | 22.4 | 20.3 | 21.4 | 20.9 | 21.5271 | 20.56 | 21.04 | 20.43 | 20.988 | 21.35 | 22.23 | 26.4134 | 19.83 | 19.34 |  | 16
| 1 | 4 | 24.6 | 20.5133 | 21.67 | 26.2 | 19.2 | 23.2 | 21.6 | 20.44 | 23.61 | 21.47 | 21 | 22.2 | 21.1 | 20.8 | 21 | 20.6037 | 20.81 | 21.44 | 20.57 | 21.2 | 21.6415 | 22.2 | 25.0203 | 20.07 | 18.87 |  | 21.3324
| 1 | 5 | 24.3 | 22.1108 | 21.75 | 26.3 | 20.4 | 22.6 | 21.4 | 20.71 | 23.69 | 21.61 | 21.2 | 22.6 | 21.1 | 21.3 | 20.5 | 21.0046 | 20.75 | 21.44 | 20.74 | 20.935 | 21.4095 | 22.25 | 25.3161 | 20.19 | 19.05 |  | 109
| 1 | 6 | 23.9 | 21.8343 | 21.88 | 25 | 20.1 | 22.2 | 21.4 | 21.08 | 23.29 | 22.23 | 21.8 | 21.5 | 21.1 | 21.2 | 21.7 | 21.0963 | 21.22 | 20.96 | 21 | 19.714 | 21.2565 | 22.39 | 25.9312 | 20.26 | 18.69
| Mean |  | 24.1 | 21.4882 | 21.7683 | 25.95 | 20.25 | 22.6167 | 21.4 | 20.7183 | 22.9267 | 21.7433 | 21.2333 | 22.2 | 21.1 | 21.1167 | 21.0667 | 21.1041 | 20.8633 | 21.1867 | 20.6533 | 20.6372 | 21.5501 | 22.2767 | 25.7538 | 20.08 | 19.02
| Median |  | 24.05 | 21.7717 | 21.79 | 26.2 | 20.15 | 22.6 | 21.4 | 20.645 | 22.955 | 21.685 | 21.25 | 22.3 | 21.1 | 21.25 | 20.95 | 21.1402 | 20.84 | 21.24 | 20.635 | 20.807 | 21.5255 | 22.24 | 25.921 | 20.115 | 18.96
| Std Dev. |  | 0.31623 | 0.65032 | 0.14148 | 0.61237 | 0.73959 | 0.392 | 0.38471 | 0.36641 | 0.69968 | 0.29824 | 0.41312 | 0.65422 | 0.57966 | 0.38687 | 0.48442 | 0.30221 | 0.22196 | 0.32562 | 0.2082 | 0.54618 | 0.25372 | 0.13471 | 0.50041 | 0.15849 | 0.27698
| Rel.Std.Dev. |  | 1.312% | 3.026% | 0.6499% | 2.36% | 3.652% | 1.733% | 1.798% | 1.769% | 3.052% | 1.372% | 1.946% | 2.947% | 2.747% | 1.832% | 2.299% | 1.432% | 1.064% | 1.537% | 1.008% | 2.647% | 1.177% | 0.6047% | 1.943% | 0.7893% | 1.456%
| PDM3 |  | 12.97% | 0.7304% | 2.044% | 21.65% | -5.074% | 6.02% | 0.3171% | -2.878% | 7.474% | 1.927% | -0.4642% | 4.067% | -1.089% | -1.011% | -1.245% | -1.07% | -2.199% | -0.683% | -3.183% | -3.259% | 1.021% | 4.427% | 20.73% | -5.871% | -10.84%
| Z-Score (Absolute) |  | 3.61 | 0.32 | 0.67 | 5.94 | 1.24 | 1.74 | 0.21 | 0.65 | 2.13 | 0.64 | 0 | 1.22 | 0.17 | 0.15 | 0.21 | 0.16 | 0.47 | 0.06 | 0.73 | 0.75 | 0.4 | 1.32 | 5.7 | 1.45 | 2.79
| Analytical results for Gd in OREAS 471 (Certified Value 8.2 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 8.3 | 7.5839 | 9.95 | 8.92 | 8.2943 | 8.03 | 7.9 | 8.05 | 9.5 | 8.47 | 7.67 | 7.55 | 6.85 | 6.92 | 6.9 | 7.756 | 8.63 | 8 | 8.3 | 7.43 | 8.6227 | 8.3238 | 8.9542 | 8.85 | 8.46 |  | 1
| 1 | 2 | 7.6 | 7.9033 | 9.58 | 8.79 | 9.0802 | 8.15 | 7.8 | 7.55 | 9.19 | 8.17 | 7.58 | 8.19 | 6.56 | 7.19 | 7.26 | 7.9634 | 8.11 | 7.3 | 8.5 | 7.911 | 8.803 | 8.3452 | 9.1471 | 8.56 | 8.5 |  | #N/A
| 1 | 3 | 8.1 | 7.7273 | 8.92 | 9.05 | 9.1075 | 8 | 7.6 | 7.99 | 8.95 | 8.23 | 7.51 | 8.13 | 6.38 | 7.25 | 6.86 | 8.0763 | 8.23 | 6.8 | 8.2 | 8.233 | 9.1368 | 8.4318 | 9.0462 | 8.74 | 8.32 |  | 16
| 1 | 4 | 7.7 | 7.8165 | 9.74 | 8.86 | 7.9922 | 8.09 | 7.8 | 8.09 | 9.09 | 8.4 | 7.83 | 8.18 | 6.64 | 6.9 | 6.82 | 7.6585 | 8.31 | 7.3 | 8.5 | 8.393 | 9.0712 | 8.2718 | 8.8548 | 8.68 | 8.43 |  | 8.2033
| 1 | 5 | 7.4 | 8.0161 | 8.65 | 9.09 | 8.8961 | 8.16 | 7.9 | 8.1 | 9.4 | 8.48 | 7.54 | 7.93 | 6.74 | 6.99 | 6.38 | 7.5725 | 8.45 | 7.1 | 8.4 | 8.647 | 9.1506 | 8.329 | 8.9684 | 8.57 | 8.45 |  | 110
| 1 | 6 | 8 | 7.9518 | 9.05 | 8.85 | 8.6259 | 8.11 | 7.7 | 8.12 | 9.47 | 8.36 | 7.75 | 7.99 | 6.45 | 6.88 | 6.9 | 7.7177 | 8.37 | 6.9 | 8.4 | 7.272 | 9.7121 | 8.3459 | 9.1885 | 8.7 | 8.39
| Mean |  | 7.85 | 7.8331 | 9.315 | 8.9267 | 8.666 | 8.09 | 7.7833 | 7.9833 | 9.2667 | 8.3517 | 7.6467 | 7.995 | 6.6033 | 7.0217 | 6.8533 | 7.7908 | 8.35 | 7.2333 | 8.3833 | 7.981 | 9.0827 | 8.3412 | 9.0265 | 8.6833 | 8.425
| Median |  | 7.85 | 7.8599 | 9.315 | 8.89 | 8.761 | 8.1 | 7.8 | 8.07 | 9.295 | 8.38 | 7.625 | 8.06 | 6.6 | 6.955 | 6.88 | 7.7368 | 8.34 | 7.2 | 8.4 | 8.072 | 9.104 | 8.3371 | 9.0073 | 8.69 | 8.44
| Std Dev. |  | 0.33912 | 0.15879 | 0.51431 | 0.11911 | 0.45016 | 0.064187 | 0.1169 | 0.21722 | 0.22402 | 0.12703 | 0.12596 | 0.24197 | 0.17671 | 0.15917 | 0.28105 | 0.19133 | 0.18022 | 0.4274 | 0.1169 | 0.54552 | 0.37247 | 0.051973 | 0.12591 | 0.10893 | 0.062849
| Rel.Std.Dev. |  | 4.32% | 2.027% | 5.521% | 1.334% | 5.195% | 0.7934% | 1.502% | 2.721% | 2.418% | 1.521% | 1.647% | 3.027% | 2.676% | 2.267% | 4.101% | 2.456% | 2.158% | 5.909% | 1.394% | 6.835% | 4.101% | 0.6231% | 1.395% | 1.255% | 0.746%
| PDM3 |  | -4.307% | -4.513% | 13.55% | 8.818% | 5.64% | -1.381% | -5.12% | -2.682% | 12.96% | 1.808% | -6.786% | -2.54% | -19.5% | -14.4% | -16.46% | -5.029% | 1.788% | -11.82% | 2.194% | -2.71% | 10.72% | 1.681% | 10.04% | 5.851% | 2.702%
| Z-Score (Absolute) |  | 0.48 | 0.52 | 2.47 | 1.68 | 1.16 | 0 | 0.62 | 0.21 | 2.37 | 0.53 | 0.89 | 0.19 | 2.99 | 2.15 | 2.49 | 0.6 | 0.52 | 1.72 | 0.59 | 0.22 | 2 | 0.51 | 1.89 | 1.19 | 0.67
| Analytical results for Ge in OREAS 471 (Indicative Value 0.27 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 24 | 27 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [8 empty] | 2
2
| 1 | 1 | 2 | 0.28 | 0.59 | < 0.1 | 0.47 | 1.2 | 1.3 | 1.2 | 0.28 | 0.24 | 0.32 | 0.28 | 0.19 | 1.4 | < 0.05 | 0.311 | 0.1366 | 0.15 | [8 empty] | 1
| 1 | 2 | 1.8 | 0.29 | 0.58 | < 0.1 | 0.48 | 1.2 | 1.2 | 1.2 | 0.22 | 0.27 | 0.29 | 0.25 | 0.21 | 1.3 | 0.11 | 0.372 | 0.14358 | 0.16 | [8 empty] | 18
| 1 | 3 | 1.8 | 0.26 | 0.56 | 0.1 | 0.47 | 1.3 | 1.2 | 1.2 | 0.22 | 0.26 | 0.3 | 0.26 | 0.19 | 1.2 | 0.23 | 0.552 | 0.12961 | 0.16 | [8 empty] | 16
| 1 | 4 | 1.7 | 0.29 | 0.6 | < 0.1 | 0.48 | 1.3 | 1.3 | 1.2 | 0.25 | 0.27 | 0.28 | 0.28 | 0.2 | 1.4 | 0.18 | 0.48 | 0.12663 | 0.16 | [8 empty] | 0.27442
| 1 | 5 | 1.8 | 0.26 | 0.59 | < 0.1 | 0.48 | 1.2 | 1.2 | 1.1 | 0.23 | 0.26 | 0.32 | 0.28 | 0.17 | 1.3 | 0.14 | 0.384 | 0.13363 | 0.16 | [8 empty] | 30
| 1 | 6 | 2 | 0.27 | 0.54 | < 0.1 | 0.45 | 1.2 | 1.3 | 1.2 | 0.27 | 0.24 | 0.3 | 0.29 | 0.21 | 1.4 | 0.24 | 0.304 | 0.1318 | 0.16
| Mean |  | 1.85 | 0.275 | 0.57667 | 0.1 | 0.47167 | 1.2333 | 1.25 | 1.1833 | 0.245 | 0.25667 | 0.30167 | 0.27333 | 0.195 | 1.3333 | 0.18 | 0.4005 | 0.13364 | 0.15833
| Median |  | 1.8 | 0.275 | 0.585 | 0.1 | 0.475 | 1.2 | 1.25 | 1.2 | 0.24 | 0.26 | 0.3 | 0.28 | 0.195 | 1.35 | 0.18 | 0.378 | 0.13272 | 0.16
| Std Dev. |  | 0.12247 | 0.013784 | 0.022509 |  | 0.01169 | 0.05164 | 0.054772 | 0.040825 | 0.025884 | 0.013663 | 0.016021 | 0.015055 | 0.015166 | 0.08165 | 0.056125 | 0.097652 | 0.0059392 | 0.0040825
| Rel.Std.Dev. |  | 6.62% | 5.012% | 3.903% |  | 2.479% | 4.187% | 4.382% | 3.45% | 10.57% | 5.323% | 5.311% | 5.508% | 7.777% | 6.124% | 31.18% | 24.38% | 4.444% | 2.578%
| PDM3 |  | 574.1% | 0.211% | 110.1% | -63.56% | 71.88% | 349.4% | 355.5% | 331.2% | -10.72% | -6.47% | 9.928% | -0.3964% | -28.94% | 385.9% | -34.41% | 45.94% | -51.3% | -42.3%
| Z-Score (Absolute) |  | 7.29 | 0.06 | 1.35 | 1.07 | 0.86 | 4.41 | 4.49 | 4.18 | 0.2 | 0.15 | 0.06 | 0.07 | 0.44 | 4.88 | 0.63 | 0.52 | 0.72 | 0.61
| Analytical results for Hf in OREAS 471 (Certified Value 4.28 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | 27 | 29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 3.7 | 3.9416 | 3.54 | 4.1 | 4.05 | 6.95 | 4.5 | 4.47 | 3.3 | 4.78 | 4.3 | 4.2 | 4.2 | 4.2 | 4.1 | 7.9381 | 4.48 | 3.9 | 4.55 | 4.139 | 4.43 | 4.6351 | 4.6 | 5.26 |  |  | 1
| 1 | 2 | 4 | 4.0456 | 3.41 | 4.1 | 4.25 | 7.09 | 4.9 | 4.29 | 3.62 | 4.49 | 4.3 | 4.4 | 4.1 | 4.2 | 4.3 | 8.0109 | 4.55 | 3.98 | 4.5 | 4.295 | 4.59 | 4.7719 | 4.4 | 5.53 |  |  | #N/A
| 1 | 3 | 3.6 | 4.0385 | 3.21 | 4.1 | 4.11 | 7.22 | 4.9 | 4.53 | 3.61 | 4.44 | 4.2 | 4.4 | 3.9 | 4.1 | 4.2 | 8.113 | 4.4 | 3.94 | 4.59 | 4.378 | 4.56 | 4.7615 | 4.4 | 5.55 |  |  | 16
| 1 | 4 | 4.1 | 4.0253 | 3.46 | 4.1 | 3.86 | 7.47 | 4.8 | 4.4 | 3.79 | 4.88 | 4.3 | 4.5 | 4.2 | 4.2 | 4.2 | 7.7869 | 4.6 | 4.13 | 4.58 | 4.248 | 4.69 | 4.7273 | 4.4 | 5.27 |  |  | 4.2796
| 1 | 5 | 3.8 | 3.9669 | 3.19 | 4.2 | 4.41 | 7.17 | 4.7 | 4.43 | 3.91 | 4.34 | 4.2 | 4.2 | 4.1 | 4.1 | 4 | 7.6893 | 4.52 | 3.92 | 4.43 | 4.541 | 4.44 | 4.6241 | 4.5 | 5.22 |  |  | 111
| 1 | 6 | 4 | 4.056 | 3.29 | 4 | 4.12 | 6.9 | 4.6 | 4.45 | 3.89 | 4.32 | 4.3 | 4.3 | 4 | 4.1 | 4.1 | 7.9298 | 4.63 | 3.88 | 4.48 | 4.16 | 4.58 | 4.7435 | 4.5 | 5.45
| Mean |  | 3.8667 | 4.0123 | 3.35 | 4.1 | 4.1333 | 7.1333 | 4.7333 | 4.4283 | 3.6867 | 4.5417 | 4.2667 | 4.3333 | 4.0833 | 4.15 | 4.15 | 7.9113 | 4.53 | 3.9583 | 4.5217 | 4.2935 | 4.5483 | 4.7106 | 4.4667 | 5.38
| Median |  | 3.9 | 4.0319 | 3.35 | 4.1 | 4.115 | 7.13 | 4.75 | 4.44 | 3.705 | 4.465 | 4.3 | 4.35 | 4.1 | 4.15 | 4.15 | 7.9339 | 4.535 | 3.93 | 4.525 | 4.2715 | 4.57 | 4.7354 | 4.45 | 5.36
| Std Dev. |  | 0.19664 | 0.046748 | 0.14184 | 0.063246 | 0.18576 | 0.20598 | 0.1633 | 0.080602 | 0.22879 | 0.23413 | 0.05164 | 0.12111 | 0.1169 | 0.054772 | 0.10488 | 0.15262 | 0.083427 | 0.090866 | 0.062423 | 0.14977 | 0.098675 | 0.064646 | 0.08165 | 0.14724
| Rel.Std.Dev. |  | 5.085% | 1.165% | 4.234% | 1.543% | 4.494% | 2.888% | 3.45% | 1.82% | 6.206% | 5.155% | 1.21% | 2.795% | 2.863% | 1.32% | 2.527% | 1.929% | 1.842% | 2.296% | 1.381% | 3.488% | 2.169% | 1.372% | 1.828% | 2.737%
| PDM3 |  | -9.649% | -6.245% | -21.72% | -4.197% | -3.418% | 66.68% | 10.6% | 3.475% | -13.85% | 6.124% | -0.3023% | 1.255% | -4.586% | -3.028% | -3.028% | 84.86% | 5.851% | -7.507% | 5.656% | 0.3247% | 6.279% | 10.07% | 4.371% | 25.71%
| Z-Score (Absolute) |  | 1.31 | 0.89 | 2.83 | 0.63 | 0.53 | 8.3 | 1.24 | 0.34 | 1.84 | 0.67 | 0.14 | 0.06 | 0.68 | 0.48 | 0.48 | 10.59 | 0.64 | 1.04 | 0.61 | 0.06 | 0.69 | 1.17 | 0.45 | 3.14
| Analytical results for Hg in OREAS 471 (Indicative Value 0.013 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 02 | 03 | [24 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | [24 empty] | 3
3
| 1 | 1 | < 1 | 0.01 | [24 empty] | 1
| 1 | 2 | < 1 | 0.01 | [24 empty] | 5
| 1 | 3 | < 1 | 0.02 | [24 empty] | 16
| 1 | 4 | < 1 | < 0.01 | [24 empty] | 0.0125
| 1 | 5 | < 1 | 0.01 | [24 empty] | 31
| 1 | 6 | < 1 | 0.02
| Mean |  |  | 0.014
| Median |  |  | 0.01
| Std Dev. |  |  | 0.0054772
| Rel.Std.Dev. |  |  | 39.12%
| PDM3 |  |  | 12%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 471 (Certified Value 0.85 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.9 | 0.77666 | 0.8 | 0.853 | 0.91247 | 0.88 | 0.8 | 0.88 | 0.8 | 0.89 | 0.87 | 0.83 | 0.75 | 0.75 | 0.79 | 0.87713 | 0.86 | 0.8 | 0.8 | 0.757 | 0.89474 | 0.81524 | 0.84803 | 0.95 | 0.85 |  | 1
| 1 | 2 | 0.8 | 0.7806 | 0.76 | 0.855 | 0.98845 | 0.88 | 0.8 | 0.85 | 0.84 | 0.86 | 0.86 | 0.88 | 0.72 | 0.77 | 0.84 | 0.91107 | 0.85 | 0.8 | 0.9 | 0.828 | 0.82462 | 0.79173 | 0.86449 | 0.89 | 0.87 |  | #N/A
| 1 | 3 | 0.8 | 0.77358 | 0.73 | 0.861 | 0.96739 | 0.86 | 0.8 | 0.85 | 0.85 | 0.87 | 0.86 | 0.9 | 0.7 | 0.78 | 0.82 | 0.89655 | 0.84 | 0.7 | 0.9 | 0.859 | 1.1044 | 0.75253 | 0.86724 | 0.9 | 0.84 |  | 16
| 1 | 4 | 0.9 | 0.78795 | 0.78 | 0.862 | 0.86572 | 0.91 | 0.8 | 0.87 | 0.89 | 0.86 | 0.87 | 0.89 | 0.73 | 0.75 | 0.8 | 0.8999 | 0.84 | 0.8 | 0.9 | 0.877 | 0.92159 | 0.75927 | 0.83743 | 0.88 | 0.84 |  | 0.84853
| 1 | 5 | 0.8 | 0.80475 | 0.72 | 0.872 | 0.96771 | 0.9 | 0.8 | 0.89 | 0.89 | 0.87 | 0.86 | 0.86 | 0.74 | 0.75 | 0.75 | 0.88075 | 0.86 | 0.8 | 0.8 | 0.911 | 1.1168 | 0.77267 | 0.87266 | 0.91 | 0.86 |  | 112
| 1 | 6 | 0.9 | 0.8098 | 0.75 | 0.85 | 0.92401 | 0.91 | 0.8 | 0.89 | 0.92 | 0.89 | 0.88 | 0.86 | 0.72 | 0.75 | 0.8 | 0.89273 | 0.86 | 0.7 | 0.9 | 0.756 | 0.86817 | 0.78799 | 0.863 | 0.9 | 0.84
| Mean |  | 0.85 | 0.78889 | 0.75667 | 0.85883 | 0.93762 | 0.89 | 0.8 | 0.87167 | 0.865 | 0.87333 | 0.86667 | 0.87 | 0.72667 | 0.75833 | 0.8 | 0.89302 | 0.85167 | 0.76667 | 0.86667 | 0.83133 | 0.95504 | 0.77991 | 0.85881 | 0.905 | 0.85
| Median |  | 0.85 | 0.78427 | 0.755 | 0.858 | 0.9457 | 0.89 | 0.8 | 0.875 | 0.87 | 0.87 | 0.865 | 0.87 | 0.725 | 0.75 | 0.8 | 0.89464 | 0.855 | 0.8 | 0.9 | 0.8435 | 0.90817 | 0.78033 | 0.86375 | 0.9 | 0.845
| Std Dev. |  | 0.054772 | 0.015115 | 0.030111 | 0.0079352 | 0.045526 | 0.02 | 1.2162e-16 | 0.018348 | 0.043243 | 0.013663 | 0.008165 | 0.025298 | 0.017512 | 0.013292 | 0.030332 | 0.01256 | 0.0098319 | 0.05164 | 0.05164 | 0.063889 | 0.12471 | 0.023158 | 0.013311 | 0.02429 | 0.012649
| Rel.Std.Dev. |  | 6.444% | 1.916% | 3.979% | 0.9239% | 4.855% | 2.247% | 1.52e-14% | 2.105% | 4.999% | 1.564% | 0.9421% | 2.908% | 2.41% | 1.753% | 3.791% | 1.406% | 1.154% | 6.736% | 5.958% | 7.685% | 13.06% | 2.969% | 1.55% | 2.684% | 1.488%
| PDM3 |  | 0.1727% | -7.029% | -10.83% | 1.214% | 10.5% | 4.887% | -5.72% | 2.726% | 1.94% | 2.923% | 2.137% | 2.53% | -14.36% | -10.63% | -5.72% | 5.243% | 0.3691% | -9.648% | 2.137% | -2.027% | 12.55% | -8.088% | 1.211% | 6.654% | 0.1727%
| Z-Score (Absolute) |  | NA | 1.51 | 2.21 | 0 | 1.7 | 0.67 | NA | 0.28 | 0.13 | 0.31 | 0.17 | 0.24 | 2.86 | 2.17 | 1.27 | 0.74 | 0.16 | NA | NA | 0.59 | 2.08 | 1.71 | 0 | 1 | 0.19
Results from laboratories 1, 8, 19 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 471 (Certified Value 0.059 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | < 0.1 | 0.07 | 0.053 | 0.05 | 0.15 | 0.06 | 0.06 | 0.053 | 0.07 | 0.058 | 0.061 | 0.06 | 0.06 | 0.056 | 0.05879 | 0.06 | 0.07 | 0.06 | < 0.1 | < 0.08 | < 0.1 | 0.099 |  |  |  |  | 1
| 1 | 2 | < 0.1 | 0.079 | 0.056 | 0.06 | 0.15 | 0.06 | 0.06 | 0.053 | 0.06 | 0.062 | 0.065 | 0.058 | 0.061 | 0.062 | 0.05777 | 0.06 | 0.06 | 0.06 | < 0.1 | < 0.08 | < 0.1 | 0.092 |  |  |  |  | 7
| 1 | 3 | < 0.1 | 0.069 | 0.055 | 0.06 | 0.16 | 0.06 | 0.05 | 0.052 | 0.05 | 0.057 | 0.068 | 0.056 | 0.062 | 0.055 | 0.05877 | 0.05 | 0.06 | 0.06 | < 0.1 | < 0.08 | < 0.1 | 0.093 |  |  |  |  | 16
| 1 | 4 | < 0.1 | 0.074 | 0.053 | 0.05 | 0.17 | 0.06 | 0.06 | 0.056 | 0.06 | 0.065 | 0.065 | 0.061 | 0.059 | 0.057 | 0.05868 | 0.07 | 0.07 | 0.06 | < 0.1 | < 0.08 | < 0.1 | 0.093 |  |  |  |  | 0.058768
| 1 | 5 | < 0.1 | 0.063 | 0.057 | 0.05 | 0.17 | 0.06 | 0.06 | 0.057 | 0.06 | 0.063 | 0.062 | 0.057 | 0.062 | 0.051 | 0.05779 | 0.06 | 0.06 | 0.06 | < 0.1 | < 0.08 | < 0.1 | 0.088 |  |  |  |  | 113
| 1 | 6 | < 0.1 | 0.071 | 0.056 | 0.06 | 0.17 | 0.06 | 0.06 | 0.048 | 0.05 | 0.063 | 0.061 | 0.059 | 0.058 | 0.058 | 0.05633 | 0.06 | 0.06 | 0.06 | < 0.1 | < 0.08 | < 0.1 | 0.087
| Mean |  |  | 0.071 | 0.055 | 0.055 | 0.16167 | 0.06 | 0.058333 | 0.053167 | 0.058333 | 0.061333 | 0.063667 | 0.0585 | 0.060333 | 0.0565 | 0.058022 | 0.06 | 0.063333 | 0.06 |  |  |  | 0.092
| Median |  |  | 0.0705 | 0.0555 | 0.055 | 0.165 | 0.06 | 0.06 | 0.053 | 0.06 | 0.0625 | 0.0635 | 0.0585 | 0.0605 | 0.0565 | 0.058235 | 0.06 | 0.06 | 0.06 |  |  |  | 0.0925
| Std Dev. |  |  | 0.0053292 | 0.0016733 | 0.0054772 | 0.0098319 | 0 | 0.0040825 | 0.0031885 | 0.0075277 | 0.0031411 | 0.0028048 | 0.0018708 | 0.001633 | 0.0036194 | 0.00095525 | 0.0063246 | 0.005164 | 0 |  |  |  | 0.0042895
| Rel.Std.Dev. |  |  | 7.506% | 3.042% | 9.959% | 6.082% | 0% | 6.999% | 5.997% | 12.9% | 5.121% | 4.405% | 3.198% | 2.707% | 6.406% | 1.646% | 10.54% | 8.154% | 0% |  |  |  | 4.663%
| PDM3 |  |  | 20.81% | -6.412% | -6.412% | 175.1% | 2.096% | -0.7398% | -9.531% | -0.7398% | 4.365% | 8.335% | -0.4562% | 2.663% | -3.859% | -1.27% | 2.096% | 7.768% | 2.096% |  |  |  | 56.55%
| Z-Score (Absolute) |  | 1.66 | 2.48 | 0.67 | 0.67 | 20.38 | 0.31 | 0.02 | 1.04 | 0.02 | 0.58 | 1.04 | 0.02 | 0.38 | 0.38 | 0.08 | 0.31 | 0.97 | 0.31 | 1.66 | 3.37 | 1.66 | 6.63
| Analytical results for K in OREAS 471 (Certified Value 2.68 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.62 | 2.5704 | 2.614 | 2.11 | 2.59 | 2.51 | 2.67 | 2.6912 | 2.37 | 2.7334 | 2.79 | 2.55 | 2.77 | 2.64 | 2.7 | 2.7375 | 2.36 | 2.74 | 2.6214 | 2.7326 | 2.77 | 0.0070809 | 2.73 |  |  |  | 1
| 1 | 2 | 2.58 | 2.5769 | 2.6485 | 1.83 | 2.82 | 2.6 | 2.68 | 2.7365 | 2.42 | 2.6225 | 2.73 | 2.62 | 2.68 | 2.75 | 2.75 | 2.7076 | 2.66 | 2.76 | 2.7147 | 2.7804 | 2.79 | 0.0072994 | 2.68 |  |  |  | #N/A
| 1 | 3 | 2.59 | 2.5808 | 2.5927 | 1.79 | 2.64 | 2.59 | 2.69 | 2.632 | 2.42 | 2.6703 | 2.69 | 2.64 | 2.66 | 2.79 | 2.62 | 2.7266 | 2.61 | 2.75 | 2.8037 | 2.771 | 2.79 | 0.0074346 | 2.65 |  |  |  | 16
| 1 | 4 | 2.66 | 2.5584 | 2.6278 | 1.9 | 2.57 | 2.62 | 2.64 | 2.726 | 2.49 | 2.6517 | 2.82 | 2.69 | 2.7 | 2.64 | 2.67 | 2.7447 | 2.73 | 2.77 | 2.7549 | 2.7789 | 2.76 | 0.0076607 | 2.68 |  |  |  | 2.6795
| 1 | 5 | 2.58 | 2.6262 | 2.5968 | 1.84 | 2.7 | 2.49 | 2.6 | 2.7042 | 2.47 | 2.6538 | 2.71 | 2.63 | 2.71 | 2.7 | 2.61 | 2.7256 | 2.61 | 2.74 | 2.7168 | 2.7879 | 2.77 | 0.0073647 | 2.69 |  |  |  | 114
| 1 | 6 | 2.67 | 2.6527 | 2.6462 | 1.81 | 2.69 | 2.53 | 2.62 | 2.7217 | 2.34 | 2.6987 | 2.84 | 2.64 | 2.71 | 2.62 | 2.68 | 2.7638 | 2.31 | 2.73 | 2.5217 | 2.7764 | 2.76 | 0.0071714 | 2.66
| Mean |  | 2.6167 | 2.5942 | 2.621 | 1.88 | 2.6683 | 2.5567 | 2.65 | 2.7019 | 2.4183 | 2.6717 | 2.7633 | 2.6283 | 2.705 | 2.69 | 2.6717 | 2.7343 | 2.5467 | 2.7483 | 2.6889 | 2.7712 | 2.7733 | 0.0073353 | 2.6817
| Median |  | 2.605 | 2.5789 | 2.6209 | 1.835 | 2.665 | 2.56 | 2.655 | 2.713 | 2.42 | 2.662 | 2.76 | 2.635 | 2.705 | 2.67 | 2.675 | 2.7321 | 2.61 | 2.745 | 2.7157 | 2.7777 | 2.77 | 0.0073321 | 2.68
| Std Dev. |  | 0.040332 | 0.036801 | 0.023971 | 0.11866 | 0.090646 | 0.053541 | 0.035777 | 0.037883 | 0.057067 | 0.039189 | 0.06186 | 0.04535 | 0.037283 | 0.068702 | 0.051929 | 0.019165 | 0.17049 | 0.01472 | 0.10147 | 0.019723 | 0.013663 | 0.00020467 | 0.027869
| Rel.Std.Dev. |  | 1.541% | 1.419% | 0.9146% | 6.312% | 3.397% | 2.094% | 1.35% | 1.402% | 2.36% | 1.467% | 2.239% | 1.725% | 1.378% | 2.554% | 1.944% | 0.7009% | 6.695% | 0.5356% | 3.774% | 0.7117% | 0.4926% | 2.79% | 1.039%
| PDM3 |  | -2.343% | -3.181% | -2.182% | -29.84% | -0.4151% | -4.583% | -1.099% | 0.8389% | -9.745% | -0.2882% | 3.13% | -1.908% | 0.9534% | 0.3935% | -0.2907% | 2.047% | -4.956% | 2.571% | 0.3513% | 3.424% | 3.504% | -99.73% | 0.08254%
| Z-Score (Absolute) |  | 0.67 | 0.95 | 0.62 | 9.71 | 0.04 | 1.41 | 0.27 | 0.37 | 3.11 | 0 | 1.12 | 0.53 | 0.41 | 0.22 | 0 | 0.77 | 1.53 | 0.94 | 0.21 | 1.22 | 1.25 | 32.67 | 0.12
| Analytical results for La in OREAS 471 (Certified Value 103 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 99.5 | 97.214 | 103 | 111 | 78.7 | 100 | 99.6 | 112.87 | 122.29 | 105.43 | 109 | 97.6 | 112 | 106.5 | 106 | 100.5155 | 110 | 92.9 | 113.4 | 84.413 | 104.521 | 107.02 | 109.441 | 101.9 | 94.41 |  | 1
| 1 | 2 | 99.6 | 97.6963 | 103.7 | 106 | 87.5 | 102 | 100 | 103.14 | 128.59 | 103.07 | 105.5 | 104 | 109.5 | 112 | 107.5 | 101.5745 | 108 | 93.9 | 114.3 | 86.169 | 96.0741 | 109.13 | 111.1681 | 103 | 94.71 |  | #N/A
| 1 | 3 | 99.8 | 99.0473 | 102.1 | 111 | 96.7 | 102 | 99.3 | 111.18 | 131.72 | 105.24 | 104.5 | 103.5 | 104.5 | 112 | 108.5 | 103.4401 | 109 | 85.6 | 112.5 | 89.084 | 100.5459 | 108.2 | 111.2589 | 102.2 | 94.91 |  | 16
| 1 | 4 | 102 | 97.8291 | 103.1 | 110 | 79.2 | 101 | 98.8 | 111.82 | 134.49 | 106.36 | 110 | 102.5 | 109 | 106.5 | 106 | 100.7656 | 110 | 88.5 | 115.3 | 90.058 | 108.1606 | 109.1 | 109.1677 | 102.5 | 94.88 |  | 103.0537
| 1 | 5 | 100 | 99.8797 | 102.1 | 112 | 95.2 | 101 | 99.1 | 114.04 | 137.34 | 106.72 | 106.5 | 100.5 | 109 | 110.5 | 93.1 | 100.0906 | 109 | 89.4 | 114.4 | 91.8 | 105.5519 | 109.83 | 110.3068 | 102.7 | 94.32 |  | 115
| 1 | 6 | 103 | 99.5838 | 103.4 | 107 | 87.7 | 100 | 98.5 | 114.82 | 145.26 | 106.47 | 112 | 100 | 106 | 107 | 104.5 | 98.3845 | 111 | 88.2 | 112.9 | 82.97 | 99.2411 | 108.33 | 110.2892 | 102.7 | 94.84
| Mean |  | 100.65 | 98.5417 | 102.9 | 109.5 | 87.5 | 101 | 99.2167 | 111.3117 | 133.2817 | 105.5483 | 107.9167 | 101.35 | 108.3333 | 109.0833 | 104.2667 | 100.7951 | 109.5 | 89.75 | 113.8 | 87.4157 | 102.3491 | 108.6017 | 110.2719 | 102.5 | 94.6783
| Median |  | 99.9 | 98.4382 | 103.05 | 110.5 | 87.6 | 101 | 99.2 | 112.345 | 133.105 | 105.895 | 107.75 | 101.5 | 109 | 108.75 | 106 | 100.6406 | 109.5 | 88.95 | 113.85 | 87.6265 | 102.5335 | 108.715 | 110.298 | 102.6 | 94.775
| Std Dev. |  | 1.4775 | 1.1061 | 0.66633 | 2.429 | 7.6197 | 0.89443 | 0.54191 | 4.2245 | 7.8317 | 1.3515 | 2.8882 | 2.4321 | 2.6771 | 2.7096 | 5.6415 | 1.6727 | 1.0488 | 3.113 | 1.0507 | 3.4445 | 4.4952 | 0.97704 | 0.85888 | 0.39497 | 0.25373
| Rel.Std.Dev. |  | 1.468% | 1.122% | 0.6476% | 2.218% | 8.708% | 0.8856% | 0.5462% | 3.795% | 5.876% | 1.28% | 2.676% | 2.4% | 2.471% | 2.484% | 5.411% | 1.66% | 0.9578% | 3.469% | 0.9233% | 3.94% | 4.392% | 0.8997% | 0.7789% | 0.3853% | 0.268%
| PDM3 |  | -2.332% | -4.378% | -0.1491% | 6.255% | -15.09% | -1.993% | -3.723% | 8.013% | 29.33% | 2.421% | 4.719% | -1.653% | 5.123% | 5.851% | 1.177% | -2.192% | 6.255% | -12.91% | 10.43% | -15.17% | -0.6837% | 5.384% | 7.004% | -0.5373% | -8.127%
| Z-Score (Absolute) |  | 0.28 | 0.54 | 0 | 0.82 | 1.91 | 0.24 | 0.46 | 1.04 | 3.77 | 0.33 | 0.62 | 0.19 | 0.67 | 0.77 | 0.17 | 0.26 | 0.82 | 1.63 | 1.35 | 1.92 | 0.07 | 0.71 | 0.91 | 0.05 | 1.02
| Analytical results for Li in OREAS 471 (Certified Value 16.4 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 25 | 26 | 27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 16 | 16.9183 | 16 | 18.9 | 17 | 16.67 | 16.2 | 16.6 | 14 | 16.5 | 17.2 | 15.4 | 17.3 | 16.2 | 16.2 | 14.9394 | 16 | 15.4 | 15.7 | 16.7 | 16.2987 | 16.8 |  |  |  |  | 1
| 1 | 2 | 16 | 17.1029 | 16 | 18.9 | 18 | 16.47 | 16.4 | 16.2 | 15 | 16.4 | 17 | 15.6 | 16.4 | 16.6 | 16.8 | 15.418 | 16.2 | 16.1 | 16.1 | 16.8 | 16.4894 | 17 |  |  |  |  | #N/A
| 1 | 3 | 16 | 17.2549 | 16 | 18.1 | 18 | 15.91 | 16.5 | 16 | 15 | 17.2 | 16.2 | 16 | 16.9 | 16.7 | 15.5 | 15.1747 | 16.1 | 15.8 | 15.6 | 16.9 | 16.3649 | 16.9 |  |  |  |  | 16
| 1 | 4 | 16 | 17.3658 | 16 | 17.6 | 16 | 17.53 | 16.3 | 16.4 | 15 | 16.8 | 17 | 17 | 16.8 | 16.2 | 15.8 | 15.3666 | 16.4 | 16.7 | 15.9 | 17 | 15.4272 | 17.2 |  |  |  |  | 16.3735
| 1 | 5 | 16 | 16.9666 | 16 | 18.2 | 18 | 16.78 | 16.2 | 16.6 | 16 | 16.3 | 16.5 | 15.8 | 16.7 | 16.4 | 15.9 | 14.9222 | 16.2 | 15.8 | 15.6 | 17 | 16.0368 | 17.3 |  |  |  |  | 116
| 1 | 6 | 17 | 17.72 | 16 | 17.6 | 18 | 16.67 | 16.5 | 16.2 | 14 | 16.8 | 17.4 | 15.6 | 16.9 | 16.4 | 16.6 | 14.5233 | 16.3 | 15.4 | 15.9 | 17 | 15.8733 | 17
| Mean |  | 16.1667 | 17.2214 | 16 | 18.2167 | 17.5 | 16.6717 | 16.35 | 16.3333 | 14.8333 | 16.6667 | 16.8833 | 15.9 | 16.8333 | 16.4167 | 16.1333 | 15.0574 | 16.2 | 15.8667 | 15.8 | 16.9 | 16.0817 | 17.0333
| Median |  | 16 | 17.1789 | 16 | 18.15 | 18 | 16.67 | 16.35 | 16.3 | 15 | 16.65 | 17 | 15.7 | 16.85 | 16.4 | 16.05 | 15.057 | 16.2 | 15.8 | 15.8 | 16.95 | 16.1678 | 17
| Std Dev. |  | 0.40825 | 0.2971 | 0 | 0.58452 | 0.83666 | 0.52331 | 0.13784 | 0.24221 | 0.75277 | 0.33267 | 0.44907 | 0.57619 | 0.29439 | 0.20412 | 0.49666 | 0.33373 | 0.14142 | 0.48854 | 0.2 | 0.12649 | 0.39178 | 0.18619
| Rel.Std.Dev. |  | 2.525% | 1.725% | 0% | 3.209% | 4.781% | 3.139% | 0.8431% | 1.483% | 5.075% | 1.996% | 2.66% | 3.624% | 1.749% | 1.243% | 3.078% | 2.216% | 0.873% | 3.079% | 1.266% | 0.7485% | 2.436% | 1.093%
| PDM3 |  | -1.263% | 5.179% | -2.281% | 11.26% | 6.88% | 1.821% | -0.1433% | -0.2451% | -9.406% | 1.791% | 3.114% | -2.892% | 2.809% | 0.2639% | -1.467% | -8.038% | -1.059% | -3.095% | -3.502% | 3.216% | -1.782% | 4.03%
| Z-Score (Absolute) |  | NA | 1.21 | NA | 2.65 | NA | 0.42 | 0.05 | 0.07 | NA | 0.41 | 0.72 | 0.7 | 0.65 | 0.05 | 0.36 | 1.92 | 0.26 | 0.75 | 0.84 | 0.75 | 0.44 | 0.94
Results from laboratories 1, 3, 6 and 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 471 (Certified Value 0.3 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.3 | 0.28356 | 0.3 | 0.29 | 0.25 | 0.29 | 0.3 | 0.31 | 0.28 | 0.32 | 0.31 | 0.29 | 0.26 | 0.29 | 0.29 | 0.32001 | 0.31 | 0.3 | 0.3 | 0.292 | 0.31104 | 0.25198 | 0.29172 | 0.32 | < 0.5 |  | 1
| 1 | 2 | 0.2 | 0.29253 | 0.29 | 0.281 | 0.27 | 0.3 | 0.31 | 0.28 | 0.29 | 0.29 | 0.31 | 0.3 | 0.25 | 0.29 | 0.3 | 0.33034 | 0.3 | 0.3 | 0.3 | 0.316 | 0.28103 | 0.2264 | 0.29586 | 0.3 | < 0.5 |  | #N/A
| 1 | 3 | 0.2 | 0.27207 | 0.27 | 0.294 | 0.27 | 0.27 | 0.3 | 0.31 | 0.29 | 0.3 | 0.29 | 0.3 | 0.24 | 0.29 | 0.29 | 0.32075 | 0.31 | 0.2 | 0.3 | 0.32 | 0.35607 | 0.24289 | 0.30192 | 0.3 | < 0.5 |  | 16
| 1 | 4 | 0.2 | 0.30176 | 0.29 | 0.286 | 0.24 | 0.29 | 0.31 | 0.31 | 0.32 | 0.32 | 0.31 | 0.31 | 0.25 | 0.28 | 0.29 | 0.31615 | 0.3 | 0.3 | 0.3 | 0.325 | 0.30738 | 0.25151 | 0.29719 | 0.3 | < 0.5 |  | 0.3007
| 1 | 5 | 0.3 | 0.29156 | 0.27 | 0.302 | 0.26 | 0.31 | 0.31 | 0.31 | 0.31 | 0.3 | 0.3 | 0.31 | 0.25 | 0.29 | 0.28 | 0.32233 | 0.31 | 0.3 | 0.3 | 0.35 | 0.39391 | 0.24133 | 0.29451 | 0.31 | < 0.5 |  | 117
| 1 | 6 | 0.3 | 0.30414 | 0.28 | 0.286 | 0.25 | 0.29 | 0.31 | 0.32 | 0.33 | 0.31 | 0.32 | 0.31 | 0.25 | 0.28 | 0.28 | 0.32032 | 0.32 | 0.3 | 0.3 | 0.283 | 0.2978 | 0.26767 | 0.29693 | 0.3 | < 0.5
| Mean |  | 0.25 | 0.29094 | 0.28333 | 0.28983 | 0.25667 | 0.29167 | 0.30667 | 0.30667 | 0.30333 | 0.30667 | 0.30667 | 0.30333 | 0.25 | 0.28667 | 0.28833 | 0.32165 | 0.30833 | 0.28333 | 0.3 | 0.31433 | 0.32454 | 0.24696 | 0.29635 | 0.305
| Median |  | 0.25 | 0.29204 | 0.285 | 0.288 | 0.255 | 0.29 | 0.31 | 0.31 | 0.3 | 0.305 | 0.31 | 0.305 | 0.25 | 0.29 | 0.29 | 0.32054 | 0.31 | 0.3 | 0.3 | 0.318 | 0.30921 | 0.2472 | 0.2964 | 0.3
| Std Dev. |  | 0.054772 | 0.011873 | 0.012111 | 0.0073869 | 0.012111 | 0.013292 | 0.005164 | 0.013663 | 0.019664 | 0.012111 | 0.010328 | 0.008165 | 0.0063246 | 0.005164 | 0.0075277 | 0.0047222 | 0.0075277 | 0.040825 | 0 | 0.024089 | 0.042169 | 0.013754 | 0.0033818 | 0.0083666
| Rel.Std.Dev. |  | 21.91% | 4.081% | 4.274% | 2.549% | 4.718% | 4.557% | 1.684% | 4.455% | 6.483% | 3.949% | 3.368% | 2.692% | 2.53% | 1.801% | 2.611% | 1.468% | 2.441% | 14.41% | 0% | 7.663% | 12.99% | 5.569% | 1.141% | 2.743%
| PDM3 |  | -16.86% | -3.248% | -5.777% | -3.615% | -14.64% | -3.006% | 1.983% | 1.983% | 0.8742% | 1.983% | 1.983% | 0.8742% | -16.86% | -4.668% | -4.114% | 6.965% | 2.537% | -5.777% | -0.2343% | 4.532% | 7.925% | -17.87% | -1.447% | 1.428%
| Z-Score (Absolute) |  | NA | 0.63 | 1.18 | 0.71 | 3.08 | 0.58 | 0.49 | 0.49 | 0.25 | 0.49 | 0.49 | 0.25 | 3.55 | 0.94 | 0.82 | 1.55 | 0.61 | NA | NA | 1.03 | 1.76 | 3.77 | 0.25 | 0.37 | 3.55
Results from laboratories 1, 19 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 471 (Certified Value 0.751 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.76 | 0.71042 | 0.8057 | 0.77 | 0.66 | 0.59 | 0.754 | 0.7621 | 0.708 | 0.7877 | 0.75 | 0.7 | 0.75 | 0.71 | 0.75 | 0.7551 | 0.78 | 0.76 | 0.73715 | 0.82258 | 0.76 | 0.00029319 | 0.76 |  |  |  | 1
| 1 | 2 | 0.75 | 0.71353 | 0.8125 | 0.77 | 0.74 | 0.64 | 0.754 | 0.7982 | 0.718 | 0.7595 | 0.72 | 0.72 | 0.71 | 0.73 | 0.76 | 0.7421 | 0.77 | 0.76 | 0.759 | 0.82947 | 0.75 | 0.0001858 | 0.77 |  |  |  | #N/A
| 1 | 3 | 0.75 | 0.71471 | 0.7944 | 0.76 | 0.71 | 0.63 | 0.751 | 0.7497 | 0.723 | 0.7691 | 0.71 | 0.73 | 0.73 | 0.73 | 0.72 | 0.7446 | 0.74 | 0.75 | 0.7682 | 0.826 | 0.75 | 0.00029956 | 0.77 |  |  |  | 16
| 1 | 4 | 0.77 | 0.71622 | 0.8033 | 0.79 | 0.68 | 0.67 | 0.75 | 0.7769 | 0.735 | 0.7731 | 0.74 | 0.78 | 0.73 | 0.7 | 0.73 | 0.7588 | 0.76 | 0.76 | 0.76245 | 0.82506 | 0.76 | 0.00020978 | 0.77 |  |  |  | 0.75125
| 1 | 5 | 0.75 | 0.72552 | 0.7988 | 0.76 | 0.72 | 0.64 | 0.744 | 0.768 | 0.731 | 0.7662 | 0.71 | 0.72 | 0.73 | 0.72 | 0.7 | 0.7511 | 0.74 | 0.76 | 0.75785 | 0.82288 | 0.76 | 0.0002237 | 0.78 |  |  |  | 118
| 1 | 6 | 0.76 | 0.73319 | 0.8094 | 0.76 | 0.71 | 0.68 | 0.746 | 0.7738 | 0.688 | 0.7857 | 0.75 | 0.71 | 0.74 | 0.72 | 0.75 | 0.7529 | 0.76 | 0.76 | 0.71875 | 0.82339 | 0.76 | 0.00022426 | 0.78
| Mean |  | 0.75667 | 0.71893 | 0.80402 | 0.76833 | 0.70333 | 0.64167 | 0.74983 | 0.77145 | 0.71717 | 0.77355 | 0.73 | 0.72667 | 0.73167 | 0.71833 | 0.735 | 0.75077 | 0.75833 | 0.75833 | 0.75057 | 0.8249 | 0.75667 | 0.00023938 | 0.77167
| Median |  | 0.755 | 0.71546 | 0.8045 | 0.765 | 0.71 | 0.64 | 0.7505 | 0.7709 | 0.7205 | 0.7711 | 0.73 | 0.72 | 0.73 | 0.72 | 0.74 | 0.752 | 0.76 | 0.76 | 0.75843 | 0.82422 | 0.76 | 0.00022398 | 0.77
| Std Dev. |  | 0.008165 | 0.008645 | 0.0066931 | 0.01169 | 0.028752 | 0.031885 | 0.0041191 | 0.016268 | 0.017198 | 0.011129 | 0.018974 | 0.028048 | 0.013292 | 0.01169 | 0.022583 | 0.0063421 | 0.016021 | 0.0040825 | 0.018808 | 0.002604 | 0.005164 | 4.6344e-05 | 0.0075277
| Rel.Std.Dev. |  | 1.079% | 1.202% | 0.8325% | 1.522% | 4.088% | 4.969% | 0.5493% | 2.109% | 2.398% | 1.439% | 2.599% | 3.86% | 1.817% | 1.627% | 3.073% | 0.8448% | 2.113% | 0.5383% | 2.506% | 0.3157% | 0.6825% | 19.36% | 0.9755%
| PDM3 |  | 0.7215% | -4.301% | 7.024% | 2.274% | -6.378% | -14.59% | -0.1881% | 2.689% | -4.536% | 2.969% | -2.828% | -3.272% | -2.606% | -4.381% | -2.163% | -0.06389% | 0.9433% | 0.9433% | -0.09052% | 9.804% | 0.7215% | -99.97% | 2.718%
| Z-Score (Absolute) |  | 0.2 | 1.03 | 1.74 | 0.57 | 1.54 | 3.55 | 0.03 | 0.68 | 1.09 | 0.74 | 0.67 | 0.78 | 0.62 | 1.05 | 0.51 | 0 | 0.25 | 0.25 | 0 | 2.42 | 0.2 | NA | 0.68
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Mn in OREAS 471 (Certified Value 0.038 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.0381 | 0.037053 | 0.0409 | 0.0411 | 0.03 | 0.0383 | 0.0381 | 0.039 | 0.0348 | 0.0388 | 0.0392 | 0.0364 | 0.0386 | 0.0377 | 0.0379 | 0.0382 | 0.0408 | 0.0404 | 0.0368 | 0.037791 | 0.038 | 0.011337 | 0.0388 |  |  |  | 1
| 1 | 2 | 0.0373 | 0.037183 | 0.0416 | 0.0402 | 0.04 | 0.0389 | 0.0382 | 0.0436 | 0.0349 | 0.0387 | 0.0382 | 0.0375 | 0.037 | 0.0393 | 0.0381 | 0.0381 | 0.0387 | 0.041 | 0.03795 | 0.037918 | 0.0384 | 0.011474 | 0.0394 |  |  |  | #N/A
| 1 | 3 | 0.0378 | 0.037081 | 0.0405 | 0.041 | 0.03 | 0.0387 | 0.038 | 0.0386 | 0.0353 | 0.0404 | 0.0376 | 0.0375 | 0.0387 | 0.0398 | 0.037 | 0.0384 | 0.0378 | 0.0405 | 0.03795 | 0.037686 | 0.038 | 0.011123 | 0.0389 |  |  |  | 16
| 1 | 4 | 0.0386 | 0.036949 | 0.0413 | 0.0409 | 0.03 | 0.0397 | 0.0382 | 0.0398 | 0.0359 | 0.0383 | 0.0394 | 0.0386 | 0.0385 | 0.0375 | 0.0373 | 0.0387 | 0.0386 | 0.0407 | 0.0391 | 0.038173 | 0.0382 | 0.010998 | 0.0391 |  |  |  | 0.038195
| 1 | 5 | 0.0371 | 0.03755 | 0.0406 | 0.0403 | 0.04 | 0.038 | 0.0376 | 0.0393 | 0.036 | 0.0384 | 0.0384 | 0.0371 | 0.0381 | 0.0383 | 0.0366 | 0.0378 | 0.0374 | 0.0403 | 0.03795 | 0.038241 | 0.038 | 0.011072 | 0.0392 |  |  |  | 119
| 1 | 6 | 0.0387 | 0.038099 | 0.0415 | 0.0413 | 0.04 | 0.0385 | 0.0379 | 0.0396 | 0.0346 | 0.039 | 0.0399 | 0.037 | 0.038 | 0.0373 | 0.0376 | 0.0385 | 0.0381 | 0.0406 | 0.0345 | 0.037949 | 0.0382 | 0.011395 | 0.0387
| Mean |  | 0.037933 | 0.037319 | 0.041067 | 0.0408 | 0.035 | 0.038683 | 0.038 | 0.039983 | 0.03525 | 0.038933 | 0.038783 | 0.03735 | 0.03815 | 0.038317 | 0.037417 | 0.038283 | 0.038567 | 0.040583 | 0.037375 | 0.03796 | 0.038133 | 0.011233 | 0.039017
| Median |  | 0.03795 | 0.037132 | 0.0411 | 0.04095 | 0.035 | 0.0386 | 0.03805 | 0.03945 | 0.0351 | 0.03875 | 0.0388 | 0.0373 | 0.0383 | 0.038 | 0.03745 | 0.0383 | 0.03835 | 0.04055 | 0.03795 | 0.037934 | 0.0381 | 0.01123 | 0.039
| Std Dev. |  | 0.00065929 | 0.00043462 | 0.00046762 | 0.00044721 | 0.0054772 | 0.00058793 | 0.00022804 | 0.0018225 | 0.00058907 | 0.00076333 | 0.00085888 | 0.00073417 | 0.00062849 | 0.0010245 | 0.00056362 | 0.00031885 | 0.0011978 | 0.00024833 | 0.0015852 | 0.00021441 | 0.0001633 | 0.00019393 | 0.00026394
| Rel.Std.Dev. |  | 1.738% | 1.165% | 1.139% | 1.096% | 15.65% | 1.52% | 0.6001% | 4.558% | 1.671% | 1.961% | 2.215% | 1.966% | 1.647% | 2.674% | 1.506% | 0.8329% | 3.106% | 0.6119% | 4.241% | 0.5648% | 0.4282% | 1.726% | 0.6765%
| PDM3 |  | -0.6852% | -2.293% | 7.518% | 6.82% | -8.365% | 1.278% | -0.5106% | 4.682% | -7.711% | 1.933% | 1.54% | -2.212% | -0.1179% | 0.3184% | -2.038% | 0.2312% | 0.973% | 6.253% | -2.147% | -0.6162% | -0.1615% | -70.59% | 2.151%
| Z-Score (Absolute) |  | 0.25 | 0.8 | 2.53 | 2.3 | 2.86 | 0.41 | 0.19 | 1.57 | 2.64 | 0.64 | 0.5 | 0.77 | 0.06 | 0.09 | 0.71 | 0.06 | 0.31 | 2.1 | 0.75 | 0.23 | 0.07 | NA | 0.71
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Mo in OREAS 471 (Certified Value 2.07 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2 | < 10 | 3.6 | 2.16 | 1.98 | 4.05 | 2 | 2.1 | 2.65 | 2 | 1.99 | 2 | 2.09 | 2.07 | 1.97 | 2.0307 | 2.2 | 1.91 | 2.04 | 2.103 | 1.5675 | 2.18 | 1.952 | 2.09 |  |  | 1
| 1 | 2 | 2 | < 10 | 3.6 | 2.24 | 2.12 | 4.26 | 2.1 | 2.1 | 2.59 | 2.1 | 2.06 | 2.16 | 1.97 | 2.12 | 2 | 2.1223 | 2.1 | 2.09 | 2.04 | 2.159 | 1.5186 | 2.15 | 2.2539 | 2.06 |  |  | #N/A
| 1 | 3 | 2 | < 10 | 3.8 | 2.12 | 2.06 | 4.23 | 2 | 2.1 | 2.67 | 2.1 | 2.02 | 2.1 | 1.91 | 2.16 | 1.96 | 2.0741 | 2.1 | 1.96 | 2 | 2.204 | 1.5592 | 2.13 | 2.2229 | 2.05 |  |  | 16
| 1 | 4 | 2 | < 10 | 3.7 | 2.18 | 1.91 | 4.39 | 2.1 | 2.1 | 2.67 | 2.1 | 2 | 2.1 | 2.03 | 2.07 | 2 | 2.0572 | 2.1 | 2.07 | 2.04 | 2.15 | 1.6399 | 2.14 | 1.9824 | 1.97 |  |  | 2.0702
| 1 | 5 | 2 | < 10 | 3.6 | 2.16 | 2.07 | 4.35 | 2.1 | 2.1 | 2.75 | 2.1 | 2.08 | 2.08 | 2.02 | 2.15 | 1.94 | 1.9544 | 2.2 | 2.05 | 2.03 | 2.089 | 1.5303 | 2.12 | 1.9843 | 1.96 |  |  | 120
| 1 | 6 | 2 | < 10 | 3.6 | 2.09 | 2.04 | 4.31 | 2 | 2.1 | 2.73 | 2.1 | 2.07 | 2.2 | 1.96 | 2.07 | 1.99 | 2.0172 | 2.2 | 1.87 | 2.07 | 2.037 | 1.5883 | 2.13 | 2.2498 | 2.02
| Mean |  | 2 |  | 3.65 | 2.1583 | 2.03 | 4.265 | 2.05 | 2.1 | 2.6767 | 2.0833 | 2.0367 | 2.1067 | 1.9967 | 2.1067 | 1.9767 | 2.0427 | 2.15 | 1.9917 | 2.0367 | 2.1237 | 1.5673 | 2.1417 | 2.1075 | 2.025
| Median |  | 2 |  | 3.6 | 2.16 | 2.05 | 4.285 | 2.05 | 2.1 | 2.67 | 2.1 | 2.04 | 2.1 | 1.995 | 2.095 | 1.98 | 2.044 | 2.15 | 2.005 | 2.04 | 2.1265 | 1.5633 | 2.135 | 2.1036 | 2.035
| Std Dev. |  | 0 |  | 0.083666 | 0.051543 | 0.074297 | 0.12029 | 0.054772 | 0 | 0.057504 | 0.040825 | 0.038297 | 0.068896 | 0.06314 | 0.042269 | 0.024221 | 0.056784 | 0.054772 | 0.091305 | 0.022509 | 0.05922 | 0.04361 | 0.02137 | 0.14833 | 0.051672
| Rel.Std.Dev. |  | 0% |  | 2.292% | 2.388% | 3.66% | 2.82% | 2.672% | 0% | 2.148% | 1.96% | 1.88% | 3.27% | 3.162% | 2.006% | 1.225% | 2.78% | 2.548% | 4.584% | 1.105% | 2.789% | 2.782% | 0.9978% | 7.038% | 2.552%
| PDM3 |  | -3.392% |  | 76.31% | 4.256% | -1.943% | 106% | -0.9765% | 1.439% | 29.29% | 0.6337% | -1.621% | 1.761% | -3.553% | 1.761% | -4.519% | -1.331% | 3.854% | -3.794% | -1.621% | 2.582% | -24.29% | 3.451% | 1.803% | -2.184%
| Z-Score (Absolute) |  | NA | 30.88 | 16.5 | 0.62 | 0.75 | 23.05 | 0.53 | 0 | 6.14 | 0.18 | 0.67 | 0.07 | 1.1 | 0.07 | 1.31 | 0.61 | 0.53 | 1.15 | 0.67 | 0.25 | 5.67 | 0.44 | 0.08 | 0.8
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 471 (Certified Value 0.903 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.89 | 0.87032 | 0.8905 | 0.96 | 0.87 | 0.91 | 0.897 | 0.8935 | 0.887 | 0.9335 | 0.95 | 0.84 | 0.92 | 0.89 | 0.92 | 0.9088 | 0.894 | 0.901 | 0.809 | 0.8696 | 0.89 | 0.045904 | 1.03 |  |  |  | 1
| 1 | 2 | 0.86 | 0.87224 | 0.9074 | 0.95 | 0.95 | 0.95 | 0.894 | 0.9035 | 0.924 | 0.908 | 0.92 | 0.86 | 0.88 | 0.92 | 0.94 | 0.9115 | 0.928 | 0.908 | 0.838 | 0.86561 | 0.89 | 0.046204 | 1.04 |  |  |  | #N/A
| 1 | 3 | 0.88 | 0.87482 | 0.8798 | 0.97 | 0.88 | 0.93 | 0.866 | 0.876 | 0.919 | 0.9137 | 0.91 | 0.9 | 0.88 | 0.93 | 0.9 | 0.9014 | 0.913 | 0.912 | 0.865 | 0.87156 | 0.89 | 0.047014 | 1.05 |  |  |  | 16
| 1 | 4 | 0.9 | 0.8725 | 0.8943 | 0.97 | 0.87 | 0.95 | 0.89 | 0.9078 | 0.94 | 0.9229 | 0.94 | 0.89 | 0.9 | 0.88 | 0.9 | 0.9101 | 0.932 | 0.92 | 0.85 | 0.86948 | 0.89 | 0.046132 | 1.06 |  |  |  | 0.90258
| 1 | 5 | 0.87 | 0.8875 | 0.8827 | 0.94 | 0.91 | 0.9 | 0.867 | 0.9001 | 0.942 | 0.8992 | 0.92 | 0.88 | 0.89 | 0.9 | 0.9 | 0.904 | 0.939 | 0.911 | 0.842 | 0.86673 | 0.9 | 0.045535 | 1.02 |  |  |  | 121
| 1 | 6 | 0.9 | 0.90094 | 0.8975 | 0.92 | 0.91 | 0.92 | 0.878 | 0.9045 | 0.887 | 0.9172 | 0.95 | 0.89 | 0.9 | 0.88 | 0.92 | 0.9151 | 0.908 | 0.901 | 0.781 | 0.87286 | 0.88 | 0.046576 | 1.06
| Mean |  | 0.88333 | 0.87972 | 0.89203 | 0.95167 | 0.89833 | 0.92667 | 0.882 | 0.89757 | 0.9165 | 0.91575 | 0.93167 | 0.87667 | 0.895 | 0.9 | 0.91333 | 0.90848 | 0.919 | 0.90883 | 0.83083 | 0.86931 | 0.89 | 0.046228 | 1.0433
| Median |  | 0.885 | 0.87366 | 0.8924 | 0.955 | 0.895 | 0.925 | 0.884 | 0.9018 | 0.9215 | 0.91545 | 0.93 | 0.885 | 0.895 | 0.895 | 0.91 | 0.90945 | 0.9205 | 0.9095 | 0.84 | 0.86954 | 0.89 | 0.046168 | 1.045
| Std Dev. |  | 0.01633 | 0.012092 | 0.010102 | 0.019408 | 0.031252 | 0.020656 | 0.013638 | 0.011628 | 0.024517 | 0.011887 | 0.017224 | 0.022509 | 0.015166 | 0.020976 | 0.01633 | 0.0050173 | 0.016923 | 0.0072503 | 0.030564 | 0.0027612 | 0.0063246 | 0.0005167 | 0.01633
| Rel.Std.Dev. |  | 1.849% | 1.375% | 1.132% | 2.039% | 3.479% | 2.229% | 1.546% | 1.296% | 2.675% | 1.298% | 1.849% | 2.568% | 1.694% | 2.331% | 1.788% | 0.5523% | 1.841% | 0.7978% | 3.679% | 0.3176% | 0.7106% | 1.118% | 1.565%
| PDM3 |  | -2.132% | -2.533% | -1.169% | 5.438% | -0.4706% | 2.669% | -2.28% | -0.5555% | 1.542% | 1.459% | 3.223% | -2.871% | -0.8399% | -0.2859% | 1.191% | 0.654% | 1.819% | 0.6927% | -7.949% | -3.687% | -1.394% | -94.88% | 15.59%
| Z-Score (Absolute) |  | 0.63 | 0.78 | 0.29 | 2.1 | 0.03 | 1.1 | 0.69 | 0.06 | 0.69 | 0.66 | 1.3 | 0.9 | 0.17 | 0.03 | 0.57 | 0.37 | 0.79 | 0.39 | 2.73 | 1.19 | 0.37 | NA | 5.76
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Nb in OREAS 471 (Certified Value 12.1 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 11.5 | 10.5862 | 10.71 | 11.4 | 12.4 | 22.7 | 11 | 12.22 | 11.5 | 13.36 | 13 | 12.5 | 13.2 | 11.8 | 12.4 | 13.0096 | 12.2 | 11.16 | 12.64 | 11.865 | 11.2126 | 12.6 | 13.0797 | 11.3 | 11.55 |  | 1
| 1 | 2 | 11.2 | 10.7005 | 11.07 | 11.4 | 13 | 23 | 11.3 | 11.82 | 11.6 | 11.65 | 13.1 | 13.3 | 13 | 12.1 | 12.8 | 13.0604 | 12.54 | 12.14 | 12.25 | 12.191 | 11.5629 | 12.8 | 12.6211 | 11.9 | 11.93 |  | #N/A
| 1 | 3 | 11.6 | 10.8073 | 9.82 | 10.9 | 12 | 24.8 | 11.3 | 12.69 | 11.6 | 12.75 | 12.6 | 13 | 11.9 | 12 | 12.7 | 13.4224 | 11.85 | 12.28 | 12.7 | 12.939 | 11.4742 | 12.8 | 12.4243 | 11.4 | 11.55 |  | 16
| 1 | 4 | 11.5 | 10.6655 | 10.64 | 11.1 | 11.9 | 24.7 | 11.6 | 12.22 | 12.1 | 12.69 | 13 | 12.7 | 12.7 | 11.6 | 12.6 | 12.8086 | 13.52 | 12.32 | 12.43 | 12.042 | 11.3867 | 12.8 | 12.614 | 11.8 | 11.66 |  | 12.1126
| 1 | 5 | 11.4 | 10.7507 | 9.89 | 11.3 | 12.7 | 22.6 | 11.1 | 12.41 | 13.3 | 11.51 | 12.7 | 12.8 | 12.2 | 12 | 12 | 12.8884 | 12 | 11.82 | 12.53 | 10.479 | 11.5795 | 12.8 | 12.3874 | 11.8 | 11.77 |  | 122
| 1 | 6 | 11.8 | 11.2304 | 10.22 | 10.6 | 12.4 | 23.5 | 11.2 | 12.61 | 14 | 11.73 | 13.5 | 13 | 12.1 | 12 | 12.7 | 12.8284 | 12.4 | 11.68 | 12.55 | 11.901 | 11.1989 | 12.8 | 12.481 | 11.6 | 11.83
| Mean |  | 11.5 | 10.7901 | 10.3917 | 11.1167 | 12.4 | 23.55 | 11.25 | 12.3283 | 12.35 | 12.2817 | 12.9833 | 12.8833 | 12.5167 | 11.9167 | 12.5333 | 13.003 | 12.4183 | 11.9 | 12.5167 | 11.9028 | 11.4025 | 12.7667 | 12.6013 | 11.6333 | 11.715
| Median |  | 11.5 | 10.7256 | 10.43 | 11.2 | 12.4 | 23.25 | 11.25 | 12.315 | 11.85 | 12.21 | 13 | 12.9 | 12.45 | 12 | 12.65 | 12.949 | 12.3 | 11.98 | 12.54 | 11.9715 | 11.4305 | 12.8 | 12.5475 | 11.7 | 11.715
| Std Dev. |  | 0.2 | 0.2284 | 0.4962 | 0.31885 | 0.41473 | 0.98133 | 0.20736 | 0.31581 | 1.0521 | 0.75468 | 0.31885 | 0.27869 | 0.52694 | 0.18348 | 0.29439 | 0.22829 | 0.59573 | 0.4429 | 0.16046 | 0.8006 | 0.16733 | 0.08165 | 0.25336 | 0.24221 | 0.15489
| Rel.Std.Dev. |  | 1.739% | 2.117% | 4.775% | 2.868% | 3.345% | 4.167% | 1.843% | 2.562% | 8.519% | 6.145% | 2.456% | 2.163% | 4.21% | 1.54% | 2.349% | 1.756% | 4.797% | 3.722% | 1.282% | 6.726% | 1.467% | 0.6396% | 2.011% | 2.082% | 1.322%
| PDM3 |  | -5.058% | -10.92% | -14.21% | -8.222% | 2.373% | 94.43% | -7.122% | 1.781% | 1.96% | 1.396% | 7.188% | 6.363% | 3.336% | -1.618% | 3.473% | 7.351% | 2.524% | -1.755% | 3.336% | -1.732% | -5.863% | 5.4% | 4.034% | -3.957% | -3.283%
| Z-Score (Absolute) |  | 1.27 | 2.37 | 2.98 | 1.86 | 0.11 | 17.26 | 1.66 | 0 | 0.03 | 0.07 | 1.01 | 0.85 | 0.29 | 0.63 | 0.32 | 1.04 | 0.14 | 0.66 | 0.29 | 0.65 | 1.42 | 0.67 | 0.42 | 1.07 | 0.94
| Analytical results for Nd in OREAS 471 (Certified Value 82 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 0
1
| 1 | 1 | 74.6 | 72.4371 | 85.9 | 86.1 | 86.2945 | 81 | 76.4 | 83.77 | 75 | 82.73 | 86.2 | 86.1 | 86.7 | 80.7 | 84.1 | 80.3004 | 81.04 | 63.8 | 86.9 | 67.601 | 86.0059 | 74.5344 | 80.8884 | 83.4 | 84.81 |  | 1
| 1 | 2 | 72.8 | 74.8549 | 83.66 | 83.5 | 95.1776 | 80 | 76.3 | 77.87 | 78.2 | 80.4 | 81.5 | 92.2 | 84.3 | 84.2 | 84.5 | 79.5084 | 81.38 | 67 | 88.7 | 68.986 | 83.0432 | 74.6413 | 82.5984 | 80.9 | 86.16 |  | #N/A
| 1 | 3 | 74.7 | 75.6152 | 78.69 | 86.7 | 98.7517 | 81 | 78.1 | 82.11 | 80.8 | 81.84 | 83.7 | 94.1 | 80.9 | 84.9 | 82.6 | 80.7003 | 80.98 | 59.8 | 86.7 | 83.287 | 85.2609 | 74.3987 | 82.6178 | 83 | 81.96 |  | 16
| 1 | 4 | 76.9 | 74.5264 | 84.32 | 85.5 | 86.4691 | 83 | 76.5 | 82.52 | 81.4 | 81.62 | 85.5 | 91.1 | 85.3 | 80.7 | 84.1 | 79.0034 | 81.27 | 63.9 | 88.3 | 84.921 | 87.2576 | 74.2634 | 78.5523 | 79 | 85.11 |  | 81.8633
| 1 | 5 | 75.1 | 74.5285 | 77.05 | 87.1 | 98.1645 | 82 | 76.4 | 84.56 | 84 | 83.56 | 84.1 | 88.1 | 84.7 | 83.5 | 75.7 | 78.5381 | 81.76 | 63.7 | 88.3 | 87.027 | 86.4899 | 74.1617 | 81.9457 | 79.9 | 85.48 |  | 123
| 1 | 6 | 78.9 | 76.542 | 80.09 | 85.1 | 92.8735 | 82 | 77.8 | 84.27 | 86.8 | 82.94 | 89.1 | 93.5 | 84.7 | 81.5 | 82.7 | 77.713 | 82.95 | 63.1 | 87.5 | 79.652 | 83.6326 | 74.4305 | 82.2606 | 82.5 | 80.63
| Mean |  | 75.5 | 74.7507 | 81.6183 | 85.6667 | 92.9552 | 81.5 | 76.9167 | 82.5167 | 81.0333 | 82.1817 | 85.0167 | 90.85 | 84.4333 | 82.5833 | 82.2833 | 79.294 | 81.5633 | 63.55 | 87.7333 | 78.579 | 85.2817 | 74.405 | 81.4772 | 81.45 | 84.025
| Median |  | 74.9 | 74.6917 | 81.875 | 85.8 | 94.0256 | 81.5 | 76.45 | 83.145 | 81.1 | 82.285 | 84.8 | 91.65 | 84.7 | 82.5 | 83.4 | 79.2559 | 81.325 | 63.75 | 87.9 | 81.4695 | 85.6334 | 74.4146 | 82.1032 | 81.7 | 84.96
| Std Dev. |  | 2.1175 | 1.3725 | 3.5094 | 1.2925 | 5.5157 | 1.0488 | 0.8085 | 2.4729 | 4.1635 | 1.1298 | 2.5787 | 3.1494 | 1.9253 | 1.8487 | 3.3205 | 1.1128 | 0.73421 | 2.2967 | 0.82381 | 8.336 | 1.6504 | 0.17466 | 1.568 | 1.7897 | 2.2027
| Rel.Std.Dev. |  | 2.805% | 1.836% | 4.3% | 1.509% | 5.934% | 1.287% | 1.051% | 2.997% | 5.138% | 1.375% | 3.033% | 3.467% | 2.28% | 2.239% | 4.035% | 1.403% | 0.9002% | 3.614% | 0.939% | 10.61% | 1.935% | 0.2347% | 1.924% | 2.197% | 2.621%
| PDM3 |  | -7.773% | -8.688% | -0.2992% | 4.646% | 13.55% | -0.4437% | -6.042% | 0.7982% | -1.014% | 0.389% | 3.852% | 10.98% | 3.139% | 0.8796% | 0.5131% | -3.139% | -0.3664% | -22.37% | 7.171% | -4.012% | 4.176% | -9.111% | -0.4716% | -0.5048% | 2.641%
| Z-Score (Absolute) |  | 1.47 | 1.65 | 0 | 0.97 | 2.72 | 0.03 | 1.13 | 0.22 | 0.14 | 0.13 | 0.81 | 2.21 | 0.67 | 0.23 | 0.16 | 0.56 | 0.01 | 4.33 | 1.46 | 0.73 | 0.88 | 1.73 | 0.03 | 0.04 | 0.58
| Analytical results for Ni in OREAS 471 (Certified Value 32.6 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 34 | 31.9835 | 33.7 | 31 | 30.1 | 31.7 | 33 | 33.9 | 44 | 31.9 | 33.4 | 31.6 | 34.3 | 31.4 | 31.7 | 30.4269 | 33 | 32.1 | 32.7 | 28.156 | 31.63 | 33.4 | 20.9074 | 32 |  |  | 1
| 1 | 2 | 32 | 32.6881 | 33.9 | 31.5 | 32.4 | 32.7 | 33 | 33.1 | 45 | 31.6 | 34.1 | 33.4 | 33.3 | 32.5 | 32.9 | 31.8842 | 33 | 32.9 | 32.9 | 29.306 | 33.11 | 33.8 | 21.4551 | 32.6 |  |  | #N/A
| 1 | 3 | 33 | 32.4956 | 33.5 | 31.6 | 30.1 | 32.7 | 33 | 32.6 | 43 | 33.3 | 33 | 32.5 | 31.8 | 32.6 | 32 | 31.6477 | 33 | 33.7 | 32.9 | 30.695 | 33.08 | 33.7 | 21.5624 | 32.8 |  |  | 16
| 1 | 4 | 34 | 32.4534 | 34.1 | 32.4 | 29.6 | 33 | 33 | 33.2 | 45 | 31.3 | 33.9 | 32.5 | 33.3 | 31.3 | 32.2 | 31.6537 | 33 | 34.6 | 32.9 | 32.14 | 32.75 | 33.8 | 20.5857 | 32.7 |  |  | 32.6055
| 1 | 5 | 32 | 32.7513 | 33.6 | 31.3 | 29.8 | 32.1 | 33 | 33.7 | 46 | 31.6 | 34.2 | 32.5 | 33.5 | 32 | 31.6 | 30.8325 | 33 | 32.7 | 33.2 | 31.28 | 33.14 | 33.5 | 21.1215 | 32.4 |  |  | 124
| 1 | 6 | 36 | 33.0396 | 33.8 | 31.3 | 30 | 32.2 | 33 | 33.5 | 45 | 31.9 | 34.9 | 31.8 | 33.3 | 31.7 | 32.2 | 30.7052 | 34 | 32.4 | 33 | 28.037 | 32.32 | 33.6 | 21.2777 | 33.3
| Mean |  | 33.5 | 32.5686 | 33.7667 | 31.5167 | 30.3333 | 32.4 | 33 | 33.3333 | 44.6667 | 31.9333 | 33.9167 | 32.3833 | 33.25 | 31.9167 | 32.1 | 31.1917 | 33.1667 | 33.0667 | 32.9333 | 29.9357 | 32.6717 | 33.6333 | 21.1516 | 32.6333
| Median |  | 33.5 | 32.5918 | 33.75 | 31.4 | 30.05 | 32.45 | 33 | 33.35 | 45 | 31.75 | 34 | 32.5 | 33.3 | 31.85 | 32.1 | 31.2401 | 33 | 32.8 | 32.9 | 30.0005 | 32.915 | 33.65 | 21.1996 | 32.65
| Std Dev. |  | 1.5166 | 0.35508 | 0.21602 | 0.47924 | 1.0309 | 0.48166 | 0 | 0.46762 | 1.0328 | 0.70616 | 0.66156 | 0.63692 | 0.80932 | 0.54924 | 0.46476 | 0.60853 | 0.40825 | 0.92664 | 0.1633 | 1.6977 | 0.59885 | 0.1633 | 0.36265 | 0.43205
| Rel.Std.Dev. |  | 4.527% | 1.09% | 0.6398% | 1.521% | 3.398% | 1.487% | 0% | 1.403% | 2.312% | 2.211% | 1.951% | 1.967% | 2.434% | 1.721% | 1.448% | 1.951% | 1.231% | 2.802% | 0.4958% | 5.671% | 1.833% | 0.4855% | 1.715% | 1.324%
| PDM3 |  | 2.743% | -0.1132% | 3.561% | -3.339% | -6.969% | -0.6302% | 1.21% | 2.232% | 36.99% | -2.061% | 4.021% | -0.6813% | 1.977% | -2.113% | -1.55% | -4.336% | 1.721% | 1.414% | 1.005% | -8.188% | 0.203% | 3.152% | -35.13% | 0.08539%
| Z-Score (Absolute) |  | 0.82 | 0.08 | 1.07 | 1.09 | 2.23 | 0.24 | 0.33 | 0.66 | 11.57 | 0.69 | 1.22 | 0.26 | 0.58 | 0.71 | 0.53 | 1.41 | 0.5 | 0.4 | 0.27 | 2.62 | 0.02 | 0.94 | 11.08 | 0.02
| Analytical results for P in OREAS 471 (Certified Value 0.143 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.157 | 0.13453 | 0.1441 | 0.142 | 0.129 | 0.1448 | 0.144 | 0.1395 | 0.146 | 0.1406 | 0.149 | 0.138 | 0.146 | 0.141 | 0.147 | 0.1413 | 0.144 | 0.141 | 0.1456 | 0.11234 | 0.14 | 0.001308 | 0.1464 |  |  |  | 1
| 1 | 2 | 0.147 | 0.13464 | 0.1454 | 0.144 | 0.1383 | 0.1497 | 0.144 | 0.1418 | 0.148 | 0.1385 | 0.145 | 0.142 | 0.139 | 0.144 | 0.149 | 0.1449 | 0.146 | 0.143 | 0.1482 | 0.11227 | 0.14 | 0.0012621 | 0.1478 |  |  |  | #N/A
| 1 | 3 | 0.155 | 0.13517 | 0.1415 | 0.142 | 0.1361 | 0.144 | 0.139 | 0.1368 | 0.149 | 0.1399 | 0.141 | 0.146 | 0.143 | 0.146 | 0.143 | 0.1408 | 0.143 | 0.142 | 0.1508 | 0.11139 | 0.14 | 0.0012424 | 0.147 |  |  |  | 16
| 1 | 4 | 0.158 | 0.13685 | 0.1431 | 0.146 | 0.1283 | 0.1504 | 0.141 | 0.1417 | 0.149 | 0.1389 | 0.148 | 0.153 | 0.143 | 0.14 | 0.143 | 0.1467 | 0.145 | 0.14 | 0.1521 | 0.11319 | 0.14 | 0.0012614 | 0.1478 |  |  |  | 0.14307
| 1 | 5 | 0.149 | 0.13875 | 0.1416 | 0.144 | 0.1367 | 0.1449 | 0.142 | 0.1404 | 0.149 | 0.1403 | 0.143 | 0.143 | 0.142 | 0.144 | 0.14 | 0.1439 | 0.142 | 0.141 | 0.1508 | 0.11449 | 0.14 | 0.0012387 | 0.1484 |  |  |  | 125
| 1 | 6 | 0.162 | 0.1371 | 0.1439 | 0.144 | 0.1363 | 0.1443 | 0.143 | 0.141 | 0.142 | 0.1452 | 0.151 | 0.142 | 0.143 | 0.145 | 0.145 | 0.1448 | 0.144 | 0.141 | 0.1378 | 0.11111 | 0.14 | 0.0012481 | 0.1477
| Mean |  | 0.15467 | 0.13617 | 0.14327 | 0.14367 | 0.13412 | 0.14635 | 0.14217 | 0.1402 | 0.14717 | 0.14057 | 0.14617 | 0.144 | 0.14267 | 0.14333 | 0.1445 | 0.14373 | 0.144 | 0.14133 | 0.14755 | 0.11246 | 0.14 | 0.0012601 | 0.14752
| Median |  | 0.156 | 0.13601 | 0.1435 | 0.144 | 0.1362 | 0.14485 | 0.1425 | 0.1407 | 0.1485 | 0.1401 | 0.1465 | 0.1425 | 0.143 | 0.144 | 0.144 | 0.14435 | 0.144 | 0.141 | 0.1495 | 0.1123 | 0.14 | 0.0012548 | 0.14775
| Std Dev. |  | 0.0056804 | 0.0016751 | 0.0015214 | 0.0015055 | 0.0043102 | 0.0028933 | 0.0019408 | 0.001873 | 0.0027869 | 0.0024097 | 0.0038166 | 0.005099 | 0.0022509 | 0.0023381 | 0.0032094 | 0.0022739 | 0.0014142 | 0.0010328 | 0.0053125 | 0.001238 | 0 | 2.5336e-05 | 0.00070545
| Rel.Std.Dev. |  | 3.673% | 1.23% | 1.062% | 1.048% | 3.214% | 1.977% | 1.365% | 1.336% | 1.894% | 1.714% | 2.611% | 3.541% | 1.578% | 1.631% | 2.221% | 1.582% | 0.9821% | 0.7308% | 3.6% | 1.101% | 0% | 2.011% | 0.4782%
| PDM3 |  | 8.107% | -4.818% | 0.1389% | 0.4185% | -6.257% | 2.294% | -0.6299% | -2.005% | 2.865% | -1.748% | 2.166% | 0.6515% | -0.2804% | 0.1855% | 1.001% | 0.4651% | 0.6515% | -1.212% | 3.133% | -21.39% | -2.144% | -99.12% | 3.11%
| Z-Score (Absolute) |  | 2.7 | 1.7 | 0.02 | 0.08 | 2.19 | 0.72 | 0.28 | 0.75 | 0.91 | 0.66 | 0.67 | 0.16 | 0.16 | 0 | 0.28 | 0.1 | 0.16 | 0.48 | 1 | 7.35 | 0.79 | 33.81 | 1
| Analytical results for Pb in OREAS 471 (Certified Value 32.1 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 32 | 31.1311 | 28 | 32.3 | 29.9 | 28.7 | 33 | 32.4 | 36.1 | 32.1 | 34.2 | 33.4 | 31.8 | 30.2 | 32 | 32.9057 | 34.4 | 31.55 | 35.26 | 32.944 | 32.385 | 32.4 | 31.5964 | 28.5 |  |  | 1
| 1 | 2 | 31 | 31.6306 | 24 | 32 | 31.8 | 28.3 | 33 | 32.8 | 31.1 | 31.3 | 33.6 | 35.2 | 30.4 | 31.3 | 33.5 | 32.7574 | 34.4 | 32.94 | 35.79 | 31.443 | 30.348 | 32.4 | 31.6469 | 31 |  |  | #N/A
| 1 | 3 | 31 | 31.3023 | 26 | 32.8 | 30.5 | 27.7 | 32 | 32.1 | 31.9 | 31.7 | 33.1 | 35.2 | 30.1 | 31.5 | 32.4 | 32.8234 | 33.7 | 31.51 | 34.77 | 31.718 | 30.852 | 32.8 | 31.2222 | 31.4 |  |  | 16
| 1 | 4 | 32 | 31.4786 | 28 | 32.4 | 29 | 29.9 | 33 | 32.1 | 34.6 | 32 | 33.5 | 35.5 | 31.1 | 29.9 | 31.7 | 31.9674 | 34.3 | 33.15 | 36.23 | 31.196 | 31.4505 | 32.7 | 33.6019 | 29.6 |  |  | 32.123
| 1 | 5 | 30 | 31.6106 | 26 | 32.8 | 30.6 | 29.6 | 32 | 33.2 | 33.1 | 31.4 | 34 | 33.5 | 31.3 | 30.9 | 30.8 | 31.377 | 34.3 | 31.49 | 34.53 | 33.456 | 30.1955 | 32.6 | 32.5958 | 30.2 |  |  | 126
| 1 | 6 | 33 | 32.1982 | 29 | 32.4 | 29.8 | 28.8 | 33 | 33 | 38.1 | 31.6 | 34.9 | 33.9 | 30.6 | 30.1 | 32.7 | 32.8842 | 34.2 | 31.97 | 35.88 | 32.242 | 30.306 | 32.6 | 31.1344 | 31.5
| Mean |  | 31.5 | 31.5585 | 26.8333 | 32.45 | 30.2667 | 28.8333 | 32.6667 | 32.6 | 34.15 | 31.6833 | 33.8833 | 34.45 | 30.8833 | 30.65 | 32.1833 | 32.4525 | 34.2167 | 32.1017 | 35.41 | 32.1665 | 30.9228 | 32.5833 | 31.9663 | 30.3667
| Median |  | 31.5 | 31.5446 | 27 | 32.4 | 30.2 | 28.75 | 33 | 32.6 | 33.85 | 31.65 | 33.8 | 34.55 | 30.85 | 30.55 | 32.2 | 32.7904 | 34.3 | 31.76 | 35.525 | 31.98 | 30.6 | 32.6 | 31.6216 | 30.6
| Std Dev. |  | 1.0488 | 0.36658 | 1.8348 | 0.30822 | 0.94587 | 0.81404 | 0.5164 | 0.46904 | 2.6486 | 0.31885 | 0.63061 | 0.95237 | 0.63061 | 0.67454 | 0.9196 | 0.6347 | 0.26394 | 0.75468 | 0.66996 | 0.88764 | 0.85572 | 0.16021 | 0.95456 | 1.1708
| Rel.Std.Dev. |  | 3.33% | 1.162% | 6.838% | 0.9498% | 3.125% | 2.823% | 1.581% | 1.439% | 7.756% | 1.006% | 1.861% | 2.764% | 2.042% | 2.201% | 2.857% | 1.956% | 0.7714% | 2.351% | 1.892% | 2.76% | 2.767% | 0.4917% | 2.986% | 3.855%
| PDM3 |  | -1.939% | -1.757% | -16.47% | 1.018% | -5.779% | -10.24% | 1.693% | 1.485% | 6.31% | -1.369% | 5.48% | 7.244% | -3.859% | -4.585% | 0.188% | 1.026% | 6.518% | -0.06626% | 10.23% | 0.1356% | -3.736% | 1.433% | -0.4878% | -5.467%
| Z-Score (Absolute) |  | 0.46 | 0.42 | 3.87 | 0.23 | 1.36 | 2.41 | 0.39 | 0.34 | 1.47 | 0.33 | 1.28 | 1.69 | 0.91 | 1.08 | 0.04 | 0.23 | 1.52 | 0.02 | 2.39 | 0.02 | 0.89 | 0.33 | 0.12 | 1.29
| Analytical results for Pr in OREAS 471 (Certified Value 23.1 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 22.4 | 20.6143 | 26.64 | 22.2 | 23.3216 | 23.35 | 21.3 | 23.63 | 22.67 | 23.5 | 25.6 | 24.5 | 25.7 | 23.3 | 23.5 | 22.1125 | 22.42 | 20.3 | 24.1 | 21.674 | 22.7789 | 22.3338 | 23.442 | 22.22 | 21.77 |  | 1
| 1 | 2 | 21.7 | 21.2032 | 25.88 | 22.2 | 25.6201 | 23.18 | 20.9 | 22.02 | 23.5 | 23 | 23.8 | 25.7 | 24.4 | 24.2 | 23.4 | 22.2274 | 22.35 | 20.5 | 25 | 21.966 | 26.8176 | 21.6861 | 23.7184 | 22.5 | 22 |  | #N/A
| 1 | 3 | 22.5 | 21.4705 | 24.2 | 22.4 | 27.0097 | 23.32 | 20.7 | 23.18 | 24.67 | 22.83 | 23.4 | 25.9 | 22.9 | 24.4 | 23.6 | 22.663 | 22.33 | 19.3 | 24.8 | 22.751 | 25.8183 | 21.8348 | 23.8009 | 23.18 | 20.96 |  | 16
| 1 | 4 | 22.8 | 21.3101 | 26.19 | 22.3 | 22.8616 | 23.98 | 20.8 | 23.15 | 24.88 | 22.96 | 24.5 | 25.5 | 24.4 | 23.5 | 23.3 | 22.2746 | 22.68 | 19.7 | 25 | 22.497 | 23.4168 | 22.1561 | 22.976 | 21.82 | 21.92 |  | 23.1426
| 1 | 5 | 22.2 | 21.1559 | 23.91 | 22.9 | 26.2808 | 23.39 | 20.7 | 23.91 | 25.69 | 23.48 | 23.9 | 24.7 | 24.3 | 23.8 | 21.2 | 21.8578 | 22.8 | 19.5 | 24.5 | 23.049 | 25.3266 | 21.9512 | 23.5669 | 22.36 | 21.64 |  | 127
| 1 | 6 | 22.4 | 21.785 | 25.12 | 22.3 | 24.7787 | 23.45 | 21 | 23.51 | 26.32 | 23.62 | 25.5 | 25.2 | 23.9 | 23.2 | 23.2 | 21.6316 | 22.7 | 18.9 | 24.2 | 21.528 | 25.5105 | 22.2937 | 23.533 | 22.92 | 20.98
| Mean |  | 22.3333 | 21.2565 | 25.3233 | 22.3833 | 24.9787 | 23.445 | 20.9 | 23.2333 | 24.6217 | 23.2317 | 24.45 | 25.25 | 24.2667 | 23.7333 | 23.0333 | 22.1278 | 22.5467 | 19.7 | 24.6 | 22.2442 | 24.9448 | 22.0426 | 23.5062 | 22.5 | 21.545
| Median |  | 22.4 | 21.2567 | 25.5 | 22.3 | 25.1994 | 23.37 | 20.85 | 23.345 | 24.775 | 23.24 | 24.2 | 25.35 | 24.35 | 23.65 | 23.35 | 22.17 | 22.55 | 19.6 | 24.65 | 22.2315 | 25.4186 | 22.0536 | 23.5499 | 22.43 | 21.705
| Std Dev. |  | 0.36697 | 0.38816 | 1.1041 | 0.26394 | 1.6431 | 0.27718 | 0.22804 | 0.65935 | 1.3537 | 0.33861 | 0.9225 | 0.55767 | 0.9048 | 0.48854 | 0.90921 | 0.35701 | 0.20354 | 0.60663 | 0.39497 | 0.61383 | 1.5338 | 0.26053 | 0.29033 | 0.48973 | 0.46233
| Rel.Std.Dev. |  | 1.643% | 1.826% | 4.36% | 1.179% | 6.578% | 1.182% | 1.091% | 2.838% | 5.498% | 1.458% | 3.773% | 2.209% | 3.729% | 2.058% | 3.947% | 1.613% | 0.9027% | 3.079% | 1.606% | 2.76% | 6.149% | 1.182% | 1.235% | 2.177% | 2.146%
| PDM3 |  | -3.497% | -8.15% | 9.423% | -3.281% | 7.934% | 1.307% | -9.69% | 0.3922% | 6.391% | 0.385% | 5.649% | 9.106% | 4.857% | 2.553% | -0.472% | -4.385% | -2.575% | -14.88% | 6.298% | -3.882% | 7.787% | -4.753% | 1.571% | -2.777% | -6.903%
| Z-Score (Absolute) |  | 0.55 | 1.21 | 1.28 | 0.52 | 1.07 | 0.13 | 1.42 | 0 | 0.85 | 0 | 0.74 | 1.23 | 0.63 | 0.31 | 0.12 | 0.67 | 0.42 | 2.16 | 0.84 | 0.6 | 1.05 | 0.73 | 0.17 | 0.45 | 1.03
| Analytical results for Rb in OREAS 471 (Certified Value 119 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 122 | 105.9217 | 129 | 125 | 100.8 | 117 | 115 | 123.34 | 126 | 118.18 | 123 | 127 | 126.5 | 120 | 120.5 | 100.5383 | 123.57 | 92.1 | 120.1 | 117.652 | 103.132 | 123 | 113.3685 | 117.4 |  |  | 1
| 1 | 2 | 121 | 110.3225 | 133 | 115 | 109.2 | 120 | 113 | 115.83 | 129 | 116.47 | 117.5 | 133.5 | 119.5 | 126.5 | 124 | 106.727 | 122.39 | 102.1 | 121.7 | 115.278 | 102.9527 | 122.3 | 114.7389 | 114.7 |  |  | #N/A
| 1 | 3 | 120 | 109.9896 | 118.2 | 108 | 112.3 | 119 | 112 | 120.34 | 127 | 119.06 | 117.5 | 132 | 116.5 | 126 | 121.5 | 96.9965 | 122.08 | 89.8 | 118.9 | 119.307 | 104.1051 | 122.9 | 111.2304 | 117.9 |  |  | 16
| 1 | 4 | 127 | 109.288 | 127.7 | 108 | 98 | 122 | 113 | 122.8 | 133 | 118.01 | 121.5 | 132.5 | 125 | 121.5 | 121 | 101.8279 | 124.68 | 103.2 | 120.7 | 115.115 | 102.5615 | 122.9 | 109.9824 | 111.2 |  |  | 119.3019
| 1 | 5 | 121 | 112.4799 | 118.5 | 110 | 110.4 | 119 | 110 | 123.13 | 136 | 118.02 | 116 | 127.5 | 123 | 123.5 | 104.5 | 98.5093 | 123.1 | 93.1 | 122.6 | 117.352 | 102.5484 | 122.5 | 110.723 | 115.3 |  |  | 128
| 1 | 6 | 121 | 111.9163 | 121.5 | 102 | 106.3 | 119 | 113 | 125.1 | 146 | 120.88 | 124 | 133 | 122.5 | 124.5 | 121 | 98.2968 | 124.01 | 88.8 | 120.9 | 114.531 | 102.8155 | 123.5 | 113.9491 | 117.7
| Mean |  | 122 | 109.9863 | 124.65 | 111.3333 | 106.1667 | 119.3333 | 112.6667 | 121.7567 | 132.8333 | 118.4367 | 119.9167 | 130.9167 | 122.1667 | 123.6667 | 118.75 | 100.4826 | 123.305 | 94.85 | 120.8167 | 116.5392 | 103.0192 | 122.85 | 112.332 | 115.7
| Median |  | 121 | 110.156 | 124.6 | 109 | 107.75 | 119 | 113 | 122.965 | 131 | 118.1 | 119.5 | 132.25 | 122.75 | 124 | 121 | 99.5238 | 123.335 | 92.6 | 120.8 | 116.315 | 102.8841 | 122.9 | 112.2994 | 116.35
| Std Dev. |  | 2.5298 | 2.3266 | 6.1204 | 7.8909 | 5.6613 | 1.633 | 1.633 | 3.28 | 7.4677 | 1.4593 | 3.3379 | 2.8882 | 3.656 | 2.543 | 7.0905 | 3.5108 | 0.98364 | 6.2452 | 1.2781 | 1.8553 | 0.57776 | 0.41833 | 1.9393 | 2.5745
| Rel.Std.Dev. |  | 2.074% | 2.115% | 4.91% | 7.088% | 5.332% | 1.368% | 1.449% | 2.694% | 5.622% | 1.232% | 2.784% | 2.206% | 2.993% | 2.056% | 5.971% | 3.494% | 0.7977% | 6.584% | 1.058% | 1.592% | 0.5608% | 0.3405% | 1.726% | 2.225%
| PDM3 |  | 2.262% | -7.808% | 4.483% | -6.679% | -11.01% | 0.02638% | -5.562% | 2.058% | 11.34% | -0.7252% | 0.5153% | 9.736% | 2.401% | 3.659% | -0.4626% | -15.77% | 3.355% | -20.5% | 1.27% | -2.316% | -13.65% | 2.974% | -5.842% | -3.019%
| Z-Score (Absolute) |  | 0.45 | 1.37 | 0.85 | 1.17 | 1.95 | 0.04 | 0.97 | 0.41 | 2.09 | 0.09 | 0.13 | 1.8 | 0.47 | 0.7 | 0.04 | 2.82 | 0.65 | 3.67 | 0.27 | 0.38 | 2.43 | 0.58 | 1.02 | 0.51
| Analytical results for Re in OREAS 471 (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 27 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.001 | < 0.01 | < 0.003 | 0.004 | 0.015 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.003 | < 0.002 | 0.007 | [8 empty] | 1
| 1 | 2 | 0.001 | < 0.01 | < 0.003 | 0.002 | 0.015 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.008 | [8 empty] | 3
| 1 | 3 | 0.001 | < 0.01 | < 0.003 | < 0.002 | 0.016 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.008 | [8 empty] | 16
| 1 | 4 | 0.002 | < 0.01 | < 0.003 | 0.003 | 0.016 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.007 | [8 empty] | < 0.002
| 1 | 5 | < 0.001 | < 0.01 | < 0.003 | 0.002 | 0.016 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.007 | [8 empty] | 129
| 1 | 6 | < 0.001 | < 0.01 | < 0.003 | 0.002 | 0.016 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.008
| Mean |  | 0.0013333 |  |  | 0.0026 | 0.015667 | [10 empty] | 0.003 |  | 0.0075
| Median |  | 0.001 |  |  | 0.002 | 0.016 | [10 empty] | 0.003 |  | 0.0075
| Std Dev. |  | 0.00057735 |  |  | 0.00089443 | 0.0005164 | [12 empty] | 0.00054772
| Rel.Std.Dev. |  | 43.3% |  |  | 34.4% | 3.296% | [12 empty] | 7.303%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 471 (Certified Value 0.03 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 24 | 25 | 27 |  |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.04 | 0.02866 | 0.0301 | 0.03 | 0.03 | 0.04 | 0.029 | 0.0299 | 0.017 | < 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.0316 | < 0.04 | < 0.04 | < 0.05 | 0.034819 | 0.03 | 0.03 |  |  |  |  | 1
| 1 | 2 | 0.03 | 0.030442 | 0.0292 | 0.03 | 0.03 | 0.03 | 0.029 | 0.0299 | 0.012 | < 0.05 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.0304 | < 0.04 | < 0.04 | < 0.05 | 0.035325 | 0.03 | 0.03 |  |  |  |  | 22
| 1 | 3 | 0.04 | 0.028825 | 0.0282 | 0.03 | 0.03 | 0.03 | 0.028 | 0.0294 | 0.015 | < 0.05 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.0309 | < 0.04 | < 0.04 | < 0.05 | 0.035447 | 0.03 | 0.03 |  |  |  |  | 16
| 1 | 4 | 0.04 | 0.032991 | 0.0292 | 0.03 | 0.03 | 0.04 | 0.029 | 0.0301 | 0.017 | < 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.0312 | < 0.04 | < 0.04 | < 0.05 | 0.035436 | 0.03 | 0.03 |  |  |  |  | 0.029965
| 1 | 5 | 0.03 | 0.033031 | 0.029 | 0.03 | 0.03 | 0.03 | 0.028 | 0.0299 | 0.013 | < 0.05 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.0309 | < 0.04 | < 0.04 | < 0.05 | 0.035124 | 0.03 | 0.03 |  |  |  |  | 130
| 1 | 6 | 0.04 | 0.03192 | 0.0292 | 0.03 | 0.03 | 0.03 | 0.028 | 0.0301 | 0.017 | < 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.0312 | < 0.04 | < 0.04 | < 0.05 | 0.035558 | 0.03 | 0.03
| Mean |  | 0.036667 | 0.030978 | 0.02915 | 0.03 | 0.03 | 0.033333 | 0.0285 | 0.029883 | 0.015167 |  | 0.035 | 0.03 | 0.04 | 0.03 | 0.03 | 0.031033 |  |  |  | 0.035285 | 0.03 | 0.03
| Median |  | 0.04 | 0.031181 | 0.0292 | 0.03 | 0.03 | 0.03 | 0.0285 | 0.0299 | 0.016 |  | 0.035 | 0.03 | 0.04 | 0.03 | 0.03 | 0.03105 |  |  |  | 0.035381 | 0.03 | 0.03
| Std Dev. |  | 0.005164 | 0.0019721 | 0.00060581 | 0 | 0 | 0.005164 | 0.00054772 | 0.00025626 | 0.0022286 |  | 0.0054772 | 0 | 0 | 0 | 0 | 0.00040332 |  |  |  | 0.0002715 | 0 | 0
| Rel.Std.Dev. |  | 14.08% | 6.366% | 2.078% | 0% | 0% | 15.49% | 1.922% | 0.8575% | 14.69% |  | 15.65% | 0% | 0% | 0% | 0% | 1.3% |  |  |  | 0.7694% | 0% | 0%
| PDM3 |  | 22.37% | 3.381% | -2.72% | 0.1169% | 0.1169% | 11.24% | -4.889% | -0.2725% | -49.39% |  | 16.8% | 0.1169% | 33.49% | 0.1169% | 0.1169% | 3.565% |  |  |  | 17.75% | 0.1169% | 0.1169%
| Z-Score (Absolute) |  | 3.55 | 0.52 | 0.45 | 0 | 0 | 1.77 | 0.8 | 0.06 | 7.9 | 2.66 | 2.66 | 0 | 5.32 | 0 | 0 | 0.55 | 5.32 | 5.32 | 2.66 | 2.81 | 0 | 0
| Analytical results for Sb in OREAS 471 (Certified Value 0.96 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.9 | < 5 | 0.93 | 1.14 | 1.06 | 2.16 | 0.9 | 0.97 | 1.27 | 0.95 | 0.98 | 0.97 | 0.98 | 0.96 | 0.97 | 0.94678 | 0.97 | 0.92 | 0.91 | 1.724 | 0.59824 | 0.84 | 0.89216 | 0.96 |  |  | 1
| 1 | 2 | 0.9 | < 5 | 0.97 | 1.12 | 1.12 | 1.98 | 0.9 | 0.99 | 1.03 | 0.93 | 0.99 | 1.02 | 0.95 | 1.01 | 1 | 0.9616 | 0.97 | 0.98 | 0.94 | 1.738 | 0.66562 | 0.89 | 0.91151 | 0.95 |  |  | #N/A
| 1 | 3 | 1 | < 5 | 0.9 | 1.07 | 1 | 2.07 | 0.9 | 0.92 | 1 | 0.94 | 0.94 | 1.04 | 0.89 | 1 | 0.95 | 0.97002 | 0.98 | 1.02 | 0.93 | 1.785 | 0.68016 | 0.89 | 0.92413 | 0.97 |  |  | 16
| 1 | 4 | 0.9 | < 5 | 0.96 | 1.11 | 0.98 | 2.21 | 0.9 | 0.97 | 0.99 | 0.93 | 0.96 | 1 | 0.97 | 0.95 | 0.99 | 0.94559 | 0.98 | 1.05 | 0.93 | 1.736 | 0.64567 | 0.89 | 0.93426 | 0.93 |  |  | 0.95703
| 1 | 5 | 0.9 | < 5 | 0.87 | 1.09 | 1.04 | 2.12 | 0.9 | 0.97 | 1.14 | 0.96 | 0.94 | 1 | 0.89 | 0.99 | 0.93 | 0.93021 | 0.99 | 0.98 | 0.94 | 1.69 | 0.68761 | 0.88 | 0.90522 | 0.95 |  |  | 131
| 1 | 6 | 0.9 | < 5 | 0.91 | 1.07 | 1.05 | 2.14 | 1 | 1.02 | 1.09 | 0.94 | 0.98 | 1 | 0.93 | 0.97 | 0.95 | 0.94325 | 0.99 | 0.94 | 0.95 | 1.63 | 0.66525 | 0.89 | 0.91595 | 0.95
| Mean |  | 0.91667 |  | 0.92333 | 1.1 | 1.0417 | 2.1133 | 0.91667 | 0.97333 | 1.0867 | 0.94167 | 0.965 | 1.005 | 0.935 | 0.98 | 0.965 | 0.94957 | 0.98 | 0.98167 | 0.93333 | 1.7172 | 0.65709 | 0.88 | 0.91387 | 0.95167
| Median |  | 0.9 |  | 0.92 | 1.1 | 1.045 | 2.13 | 0.9 | 0.97 | 1.06 | 0.94 | 0.97 | 1 | 0.94 | 0.98 | 0.96 | 0.94619 | 0.98 | 0.98 | 0.935 | 1.73 | 0.66543 | 0.89 | 0.91373 | 0.95
| Std Dev. |  | 0.040825 |  | 0.037771 | 0.028284 | 0.04916 | 0.079917 | 0.040825 | 0.03266 | 0.10633 | 0.01169 | 0.021679 | 0.023452 | 0.038859 | 0.023664 | 0.026646 | 0.014151 | 0.0089443 | 0.048339 | 0.013663 | 0.052484 | 0.032239 | 0.02 | 0.014676 | 0.013292
| Rel.Std.Dev. |  | 4.454% |  | 4.091% | 2.571% | 4.719% | 3.782% | 4.454% | 3.355% | 9.785% | 1.241% | 2.247% | 2.334% | 4.156% | 2.415% | 2.761% | 1.49% | 0.9127% | 4.924% | 1.464% | 3.056% | 4.906% | 2.273% | 1.606% | 1.397%
| PDM3 |  | -4.217% |  | -3.521% | 14.94% | 8.844% | 120.8% | -4.217% | 1.704% | 13.55% | -1.605% | 0.833% | 5.013% | -2.302% | 2.4% | 0.833% | -0.7788% | 2.4% | 2.575% | -2.476% | 79.43% | -31.34% | -8.049% | -4.509% | -0.5602%
| Z-Score (Absolute) |  | NA | 28.81 | 0.86 | 2.46 | 1.36 | 21.53 | NA | 0.08 | 2.21 | 0.52 | 0.08 | 0.67 | 0.64 | 0.2 | 0.08 | 0.37 | 0.2 | 0.24 | 0.67 | 14.08 | 5.87 | 1.68 | 1.04 | 0.33
Results from laboratories 1 and 8 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 471 (Certified Value 12.7 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 25 | 26 | 27 | 29 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 13 | 12.0303 | 12.24 | 13.8 | 12.2 | 12.2 | 12 | 13 | 9.9 | 12.9 | 12.8 | 13.2 | 14.2 | 11.4 | 12 | 11.581 | 12.6 | 12.6 | 12.1 | 12.317 | 14.0951 | 12.59 | 12.6 | 12.856 | 12.1 | 13.27 | 1
| 1 | 2 | 12 | 12.0603 | 12.38 | 14.1 | 12.8 | 12.6 | 12 | 12.1 | 10.4 | 12.4 | 12.9 | 13.6 | 13.4 | 11.9 | 12.4 | 11.9584 | 12.5 | 11.8 | 12.2 | 13.898 | 13.9279 | 12.74 | 12.7 | 13.1249 | 12.2 | 13.55 | #N/A
| 1 | 3 | 13 | 12.0848 | 12.21 | 13.4 | 12.3 | 12.6 | 12 | 12.8 | 10.6 | 12.8 | 12.5 | 13.6 | 12.9 | 12 | 11.8 | 11.5359 | 12.6 | 12.7 | 12.1 | 14.145 | 15.0746 | 12.72 | 12.8 | 13.0493 | 12.1 | 13.47 | 16
| 1 | 4 | 13 | 12.0099 | 12.22 | 13.9 | 11.1 | 12 | 12 | 12.9 | 10.5 | 12.9 | 12.8 | 13.8 | 13.7 | 11.5 | 12 | 11.4636 | 12.7 | 12.3 | 12.2 | 14.24 | 14.4295 | 12.89 | 12.9 | 12.932 | 12.2 | 13.29 | 12.7025
| 1 | 5 | 12 | 12.2624 | 12.26 | 13.9 | 11.9 | 11.8 | 12 | 13.1 | 10.7 | 12.6 | 12.7 | 13.5 | 13.5 | 11.8 | 11.3 | 11.253 | 12.6 | 12.6 | 12.3 | 14.056 | 14.7484 | 12.96 | 12.6 | 12.8608 | 12.1 | 13.56 | 132
| 1 | 6 | 13 | 12.4115 | 12.29 | 13.2 | 11.9 | 12.6 | 12 | 13.1 | 11.3 | 13 | 13.2 | 13.2 | 13.7 | 11.6 | 12.1 | 11.3541 | 12.8 | 12.4 | 12.2 | 11.733 | 13.7788 | 12.81 | 12.9 | 13.0375 | 12.2 | 13.58
| Mean |  | 12.6667 | 12.1432 | 12.2667 | 13.7167 | 12.0333 | 12.3 | 12 | 12.8333 | 10.5667 | 12.7667 | 12.8167 | 13.4833 | 13.5667 | 11.7 | 11.9333 | 11.5243 | 12.6333 | 12.4 | 12.1833 | 13.3982 | 14.3424 | 12.785 | 12.75 | 12.9767 | 12.15 | 13.4533
| Median |  | 13 | 12.0725 | 12.25 | 13.85 | 12.05 | 12.4 | 12 | 12.95 | 10.55 | 12.85 | 12.8 | 13.55 | 13.6 | 11.7 | 12 | 11.4997 | 12.6 | 12.5 | 12.2 | 13.977 | 14.2623 | 12.775 | 12.75 | 12.9847 | 12.15 | 13.51
| Std Dev. |  | 0.5164 | 0.15938 | 0.062503 | 0.34303 | 0.56451 | 0.35214 | 0 | 0.37771 | 0.45461 | 0.22509 | 0.23166 | 0.24014 | 0.4274 | 0.23664 | 0.36697 | 0.24415 | 0.10328 | 0.32863 | 0.075277 | 1.0854 | 0.50183 | 0.13157 | 0.13784 | 0.1104 | 0.054772 | 0.13952
| Rel.Std.Dev. |  | 4.077% | 1.313% | 0.5095% | 2.501% | 4.691% | 2.863% | 0% | 2.943% | 4.302% | 1.763% | 1.807% | 1.781% | 3.15% | 2.023% | 3.075% | 2.119% | 0.8175% | 2.65% | 0.6179% | 8.101% | 3.499% | 1.029% | 1.081% | 0.8507% | 0.4508% | 1.037%
| PDM3 |  | -0.2818% | -4.403% | -3.431% | 7.984% | -5.268% | -3.168% | -5.53% | 1.03% | -16.81% | 0.5054% | 0.899% | 6.147% | 6.803% | -7.892% | -6.055% | -9.275% | -0.5443% | -2.381% | -4.087% | 5.477% | 12.91% | 0.6497% | 0.3742% | 2.159% | -4.349% | 5.911%
| Z-Score (Absolute) |  | NA | 0.68 | 0.53 | 1.27 | 0.81 | 0.48 | NA | 0.18 | 2.63 | 0.09 | 0.15 | 0.98 | 1.08 | 1.23 | 0.94 | 1.44 | 0.07 | 0.36 | 0.63 | 0.87 | 2.04 | 0.12 | 0.07 | 0.35 | 0.67 | 0.94
Results from laboratories 1 and 8 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 471 (Indicative Value 1.69 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 25 | 26 | 27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2 | < 5 | 1.27 | < 0.5 | < 2 | 3.65 | < 1 | < 0.5 | 3 | < 0.5 | 1 | < 1 | < 1 | 1 | < 1 | < 0.5 | 0.8 | 0.6 | 2.221 | 0.4 | 2.9319 | < 1 |  |  |  |  | 1
| 1 | 2 | 1.8 | < 5 | 1.34 | < 0.5 | < 2 | 3.66 | < 1 | < 0.5 | 3 | < 0.5 | 2 | < 1 | < 1 | 1 | < 1 | < 0.5 | 1 | 0.4 | 2.331 | 0.3 | 1.858 | < 1 |  |  |  |  | 1
| 1 | 3 | 2 | < 5 | 1.06 | < 0.5 | < 2 | 3.59 | < 1 | < 0.5 | 3 | < 0.5 | 2 | < 1 | < 1 | 1 | < 1 | < 0.5 | 1 | 0.5 | 2.337 | 0.4 | 2.9956 | < 1 |  |  |  |  | 16
| 1 | 4 | 2.3 | < 5 | 1.22 | < 0.5 | < 2 | 3.65 | < 1 | < 0.5 | 3 | < 0.5 | 2 | < 1 | < 1 | < 1 | < 1 | < 0.5 | 1.2 | 0.4 | 2.305 | 0.3 | 2.0978 | < 1 |  |  |  |  | 1.6859
| 1 | 5 | 2 | < 5 | 1.13 | < 0.5 | < 2 | 3.84 | < 1 | < 0.5 | 3 | < 0.5 | 3 | < 1 | < 1 | 1 | < 1 | < 0.5 | 1 | 0.5 | 2.293 | 0.4 | 2.237 | < 1 |  |  |  |  | 30
| 1 | 6 | 2.2 | < 5 | 1.26 | < 0.5 | < 2 | 3.77 | < 1 | 0.6 | 4 | < 0.5 | 2 | < 1 | < 1 | 1 | < 1 | < 0.5 | 0.8 | 0.5 | 2.099 | 0.4 | 2.2426 | < 1
| Mean |  | 2.05 |  | 1.2133 |  |  | 3.6933 |  | 0.6 | 3.1667 |  | 2 |  |  | 1 |  |  | 0.96667 | 0.48333 | 2.2643 | 0.36667 | 2.3938
| Median |  | 2 |  | 1.24 |  |  | 3.655 |  | 0.6 | 3 |  | 2 |  |  | 1 |  |  | 1 | 0.5 | 2.299 | 0.4 | 2.2398
| Std Dev. |  | 0.17607 |  | 0.10191 |  |  | 0.092664 |  |  | 0.40825 |  | 0.63246 |  |  | 0 |  |  | 0.15055 | 0.075277 | 0.091009 | 0.05164 | 0.46344
| Rel.Std.Dev. |  | 8.589% |  | 8.4% |  |  | 2.509% |  |  | 12.89% |  | 31.62% |  |  | 0% |  |  | 15.57% | 15.57% | 4.019% | 14.08% | 19.36%
| PDM3 |  | 21.59% |  | -28.03% |  |  | 119.1% |  | -64.41% | 87.83% |  | 18.63% |  |  | -40.69% |  |  | -42.66% | -71.33% | 34.31% | -78.25% | 41.99%
| Z-Score (Absolute) |  | 2.11 | 2.82 | 0.79 | 0.72 | 0.46 | 4.69 | 0.33 | 0.63 | 3.86 | 0.72 | 2.03 | 0.33 | 0.33 | 0.33 | 0.33 | 0.72 | 0.41 | 0.35 | 2.44 | 0.54 | 2.65 | 0.33
| Analytical results for Sm in OREAS 471 (Certified Value 12.3 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 11.9 | 11.3519 | 12.99 | 13.2 | 12.9992 | 12.83 | 11.6 | 12.75 | 12.34 | 12.6 | 12.8 | 11.8 | 11.95 | 11.3 | 11.5 | 12.0932 | 12.3 | 11.1 | 12.7 | 12.32 | 12.0918 | 11.1823 | 12.6274 | 13.04 | 11.99 |  | 1
| 1 | 2 | 11.3 | 11.6038 | 12.67 | 12.8 | 14.0246 | 13.09 | 11.4 | 11.72 | 12.93 | 12.16 | 12.65 | 12.3 | 11.25 | 11.8 | 11.85 | 12.0022 | 12.19 | 11.1 | 13 | 12.745 | 12.4304 | 10.651 | 12.9369 | 12.66 | 12.03 |  | #N/A
| 1 | 3 | 11.6 | 11.5602 | 11.92 | 13.2 | 14.2398 | 12.8 | 11.2 | 12.51 | 13.58 | 12.23 | 12.45 | 12.45 | 10.8 | 11.95 | 11.4 | 12.3046 | 12.21 | 9.6 | 12.7 | 13.209 | 13.7344 | 10.8035 | 13.0053 | 12.84 | 11.55 |  | 16
| 1 | 4 | 11.7 | 11.6229 | 12.68 | 12.9 | 12.6684 | 13.06 | 11.4 | 12.65 | 13.54 | 12.27 | 12.75 | 12.35 | 11.45 | 11.3 | 11.4 | 11.9417 | 12.46 | 10.7 | 12.9 | 13.45 | 12.0172 | 10.4934 | 12.6358 | 12.29 | 12.01 |  | 12.2566
| 1 | 5 | 11.9 | 11.9214 | 11.61 | 13.3 | 14.0894 | 13.08 | 11.2 | 13.19 | 14.01 | 12.58 | 12.55 | 11.9 | 11.3 | 11.7 | 10.5 | 11.7979 | 12.48 | 10.8 | 12.6 | 13.792 | 13.1291 | 10.5525 | 12.804 | 12.57 | 12.11 |  | 133
| 1 | 6 | 11.7 | 11.6522 | 12.17 | 12.9 | 13.5708 | 12.71 | 11.2 | 12.75 | 14.54 | 12.6 | 13 | 12.35 | 11.15 | 11.35 | 11.5 | 11.8596 | 12.57 | 10.5 | 12.9 | 12.253 | 13.3221 | 11.1445 | 12.8827 | 12.86 | 11.76
| Mean |  | 11.6833 | 11.6187 | 12.34 | 13.05 | 13.5987 | 12.9283 | 11.3333 | 12.595 | 13.49 | 12.4067 | 12.7 | 12.1917 | 11.3167 | 11.5667 | 11.3583 | 11.9999 | 12.3683 | 10.6333 | 12.8 | 12.9615 | 12.7875 | 10.8045 | 12.8154 | 12.71 | 11.9083
| Median |  | 11.7 | 11.6134 | 12.42 | 13.05 | 13.7977 | 12.945 | 11.3 | 12.7 | 13.56 | 12.425 | 12.7 | 12.325 | 11.275 | 11.525 | 11.45 | 11.9719 | 12.38 | 10.75 | 12.8 | 12.977 | 12.7798 | 10.7272 | 12.8434 | 12.75 | 12
| Std Dev. |  | 0.22286 | 0.18308 | 0.52635 | 0.20736 | 0.64168 | 0.1675 | 0.1633 | 0.48538 | 0.77712 | 0.20762 | 0.19494 | 0.27096 | 0.37903 | 0.28577 | 0.45212 | 0.1819 | 0.1569 | 0.55737 | 0.15492 | 0.6245 | 0.70768 | 0.29739 | 0.1569 | 0.26336 | 0.21113
| Rel.Std.Dev. |  | 1.908% | 1.576% | 4.265% | 1.589% | 4.719% | 1.296% | 1.441% | 3.854% | 5.761% | 1.673% | 1.535% | 2.222% | 3.349% | 2.471% | 3.981% | 1.516% | 1.269% | 5.242% | 1.21% | 4.818% | 5.534% | 2.752% | 1.224% | 2.072% | 1.773%
| PDM3 |  | -4.678% | -5.205% | 0.6801% | 6.473% | 10.95% | 5.48% | -7.533% | 2.761% | 10.06% | 1.224% | 3.617% | -0.5302% | -7.669% | -5.629% | -7.329% | -2.095% | 0.9112% | -13.24% | 4.433% | 5.751% | 4.331% | -11.85% | 4.558% | 3.699% | -2.842%
| Z-Score (Absolute) |  | 0.82 | 0.9 | 0.03 | 0.82 | 1.48 | 0.67 | 1.25 | 0.27 | 1.35 | 0.05 | 0.4 | 0.21 | 1.27 | 0.97 | 1.22 | 0.44 | 0 | 2.09 | 0.52 | 0.71 | 0.5 | 1.88 | 0.54 | 0.41 | 0.55
| Analytical results for Sn in OREAS 471 (Certified Value 3.05 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 3 | 2.7549 | 3.17 | 3.2 | 3.4 | 4.5 | 2.9 | 3.1 | 3.5 | 3.2 | 3 | 2.9 | 3.1 | 3.1 | 3 | 3.0788 | 3.2 | 2.9 | 2.9 | 2.779 | 2.0312 | 3.1 | 3.3 |  |  |  | 1
| 1 | 2 | 3 | 2.8228 | 3.33 | 3.2 | 3.6 | 4.4 | 3 | 3 | 3.6 | 3 | 2.9 | 3 | 3 | 3.2 | 3 | 3.1548 | 3.3 | 3.1 | 2.8 | 2.897 | 2.1095 | 3.2 | 3.2 |  |  |  | #N/A
| 1 | 3 | 3 | 2.6781 | 2.95 | 3.1 | 3.4 | 4.58 | 3 | 3 | 3.7 | 3.1 | 2.9 | 3 | 2.8 | 3.2 | 3 | 3.1515 | 3.1 | 3.3 | 2.8 | 3.032 | 1.9399 | 3.1 | 3.2 |  |  |  | 16
| 1 | 4 | 3 | 2.7615 | 3.17 | 3.1 | 3.2 | 4.49 | 3 | 3.1 | 3.6 | 3 | 3 | 3 | 3 | 3 | 3 | 3.0713 | 3.3 | 3.2 | 2.8 | 3.61 | 1.9821 | 3.2 | 3 |  |  |  | 3.0466
| 1 | 5 | 3 | 2.6704 | 2.91 | 3.2 | 3.4 | 4.56 | 3 | 3.1 | 3.7 | 2.9 | 3 | 2.9 | 3 | 3.1 | 2.9 | 3.0674 | 3.2 | 3.2 | 2.8 | 2.958 | 1.9323 | 3.2 | 3.1 |  |  |  | 134
| 1 | 6 | 3 | 2.7537 | 3.06 | 3 | 3.3 | 4.86 | 3 | 3.2 | 3.7 | 3.1 | 3 | 3 | 3 | 3.1 | 2.9 | 3.1056 | 3.2 | 3.1 | 2.9 | 2.597 | 1.9623 | 3.1 | 3.2
| Mean |  | 3 | 2.7402 | 3.0983 | 3.1333 | 3.3833 | 4.565 | 2.9833 | 3.0833 | 3.6333 | 3.05 | 2.9667 | 2.9667 | 2.9833 | 3.1167 | 2.9667 | 3.1049 | 3.2167 | 3.1333 | 2.8333 | 2.9788 | 1.9929 | 3.15 | 3.1667
| Median |  | 3 | 2.7543 | 3.115 | 3.15 | 3.4 | 4.53 | 3 | 3.1 | 3.65 | 3.05 | 3 | 3 | 3 | 3.1 | 3 | 3.0922 | 3.2 | 3.15 | 2.8 | 2.9275 | 1.9722 | 3.15 | 3.2
| Std Dev. |  | 0 | 0.05727 | 0.15677 | 0.08165 | 0.13292 | 0.1577 | 0.040825 | 0.075277 | 0.08165 | 0.10488 | 0.05164 | 0.05164 | 0.098319 | 0.075277 | 0.05164 | 0.039701 | 0.075277 | 0.13663 | 0.05164 | 0.34469 | 0.067224 | 0.054772 | 0.10328
| Rel.Std.Dev. |  | 0% | 2.09% | 5.06% | 2.606% | 3.929% | 3.455% | 1.368% | 2.441% | 2.247% | 3.439% | 1.741% | 1.741% | 3.296% | 2.415% | 1.741% | 1.279% | 2.34% | 4.36% | 1.823% | 11.57% | 3.373% | 1.739% | 3.261%
| PDM3 |  | -1.531% | -10.06% | 1.697% | 2.846% | 11.05% | 49.84% | -2.078% | 1.205% | 19.26% | 0.1105% | -2.625% | -2.625% | -2.078% | 2.299% | -2.625% | 1.913% | 5.581% | 2.846% | -7.001% | -2.225% | -34.59% | 3.393% | 3.94%
| Z-Score (Absolute) |  | NA | 2.15 | 0.05 | 0.26 | 1.8 | 9.06 | 0.66 | 0.05 | 3.33 | 0.25 | 0.76 | 0.76 | 0.66 | 0.16 | 0.76 | 0.09 | 0.77 | 0.26 | 1.58 | 0.69 | 6.75 | 0.36 | 0.47
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 471 (Certified Value 442 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 432 | 429.2512 | 508 | 535 | 393.7 | 444 | 439 | 451 | 399 | 447.27 | 465 | 429 | 464 | 431 | 452 | 405.6639 | 435.91 | 422 | 437 | 452.641 | 446.77 | 441.4 | 459.0402 | 443.4 |  |  | 1
| 1 | 2 | 431 | 429.0866 | 511 | 532 | 435.6 | 448 | 437 | 433 | 403 | 432.22 | 452 | 440 | 427 | 449 | 459 | 401.2685 | 431.84 | 426 | 440 | 453.759 | 457.29 | 442.4 | 462.0367 | 450.2 |  |  | #N/A
| 1 | 3 | 432 | 428.9274 | 499 | 518 | 427.7 | 450 | 436 | 443.54 | 408 | 437.62 | 440 | 448 | 445 | 453 | 435 | 407.3365 | 424.41 | 405 | 436 | 467.742 | 451.9 | 441.4 | 470.1434 | 444.6 |  |  | 16
| 1 | 4 | 446 | 427.7804 | 507 | 522 | 384.8 | 454 | 437 | 448.24 | 414 | 441.43 | 458 | 468 | 456 | 430 | 438 | 408.8203 | 439.11 | 421 | 433 | 449.085 | 451.98 | 442.2 | 461.3214 | 448.7 |  |  | 442.4787
| 1 | 5 | 424 | 435.8084 | 503 | 514 | 415.7 | 440 | 436 | 448.76 | 416 | 438.71 | 449 | 439 | 443 | 437 | 418 | 398.9823 | 433.9 | 413 | 443 | 456.014 | 442.76 | 442.3 | 455.3472 | 448.2 |  |  | 135
| 1 | 6 | 437 | 443.2205 | 508 | 492 | 401.7 | 446 | 437 | 455.21 | 398 | 445.14 | 465 | 437 | 459 | 442 | 443 | 406.3648 | 438.19 | 420 | 438 | 433.036 | 455.67 | 442.4 | 465.7614 | 446.5
| Mean |  | 433.6667 | 432.3457 | 506 | 518.8333 | 409.8667 | 447 | 437 | 446.625 | 406.3333 | 440.3983 | 454.8333 | 443.5 | 449 | 440.3333 | 440.8333 | 404.7394 | 433.8933 | 417.8333 | 437.8333 | 452.0462 | 451.0617 | 442.0167 | 462.275 | 446.9333
| Median |  | 432 | 429.1689 | 507.5 | 520 | 408.7 | 447 | 437 | 448.5 | 405.5 | 440.07 | 455 | 439.5 | 450.5 | 439.5 | 440.5 | 406.0144 | 434.905 | 420.5 | 437.5 | 453.2 | 451.94 | 442.25 | 461.679 | 447.35
| Std Dev. |  | 7.3394 | 6.0494 | 4.2895 | 15.4197 | 19.8527 | 4.858 | 1.0954 | 7.6833 | 7.607 | 5.4447 | 9.786 | 13.4573 | 13.4907 | 9.4163 | 14.3027 | 3.797 | 5.3645 | 7.5741 | 3.4303 | 11.2743 | 5.4573 | 0.48339 | 5.167 | 2.589
| Rel.Std.Dev. |  | 1.692% | 1.399% | 0.8477% | 2.972% | 4.844% | 1.087% | 0.2507% | 1.72% | 1.872% | 1.236% | 2.152% | 3.034% | 3.005% | 2.138% | 3.244% | 0.9381% | 1.236% | 1.813% | 0.7835% | 2.494% | 1.21% | 0.1094% | 1.118% | 0.5793%
| PDM3 |  | -1.992% | -2.29% | 14.36% | 17.26% | -7.37% | 1.022% | -1.238% | 0.9371% | -8.169% | -0.4702% | 2.792% | 0.2308% | 1.474% | -0.4849% | -0.3719% | -8.529% | -1.94% | -5.57% | -1.05% | 2.162% | 1.94% | -0.1044% | 4.474% | 1.007%
| Z-Score (Absolute) |  | 0.68 | 0.8 | 5.68 | 6.81 | 2.78 | 0.49 | 0.39 | 0.46 | 3.09 | 0.09 | 1.18 | 0.18 | 0.67 | 0.1 | 0.05 | 3.23 | 0.66 | 2.07 | 0.32 | 0.93 | 0.85 | 0.05 | 1.83 | 0.48
| Analytical results for Ta in OREAS 471 (Certified Value 0.79 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 25 | 26 | 27 | 29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.8 | 0.82641 | 0.66 | 0.77 | 1.01 | 0.74 | 0.69 | 0.76 | 1.08 | 0.77 | 0.81 | 0.74 | 0.74 | 0.74 | 0.75 | 0.84842 | 1 | 0.9 | 0.8 | < 1 | 1.2 | 0.97699 | 1.13 | 0.93 |  |  | 1
| 1 | 2 | 0.8 | 0.81804 | 0.81 | 0.73 | 0.99 | 0.8 | 0.69 | 0.74 | 1.01 | 0.69 | 0.79 | 0.78 | 0.73 | 0.75 | 0.76 | 0.85625 | 1 | 1 | 0.8 | < 1 | 1.3 | 0.96488 | 1.08 | 0.94 |  |  | #N/A
| 1 | 3 | 0.8 | 0.81836 | 0.6 | 0.72 | 0.8 | 0.81 | 0.66 | 0.78 | 1.01 | 0.74 | 0.78 | 0.78 | 0.69 | 0.74 | 0.77 | 0.86821 | 1 | 0.9 | 0.8 | < 1 | 1.3 | 0.86112 | 1.07 | 0.94 |  |  | 16
| 1 | 4 | 0.9 | 0.84451 | 1 | 0.73 | 0.86 | 0.77 | 0.7 | 0.74 | 0.98 | 0.76 | 0.78 | 0.78 | 0.72 | 0.72 | 0.76 | 0.8312 | < 1 | 1 | 0.8 | < 1 | 1.2 | 0.89462 | 1.13 | 0.94 |  |  | 0.79103
| 1 | 5 | 0.8 | 0.80026 | 0.9 | 0.75 | 0.88 | 0.71 | 0.7 | 0.78 | 1.14 | 0.71 | 0.76 | 0.74 | 0.71 | 0.75 | 0.71 | 0.83252 | < 1 | 1 | 0.8 | < 1 | 1.3 | 0.88286 | 1.14 | 0.93 |  |  | 136
| 1 | 6 | 0.8 | 0.81987 | 0.61 | 0.7 | 0.95 | 0.75 | 0.69 | 0.8 | 1.15 | 0.75 | 0.81 | 0.78 | 0.72 | 0.74 | 0.74 | 0.87425 | < 1 | 0.9 | 0.8 | < 1 | 1.2 | 0.91602 | 1.16 | 0.92
| Mean |  | 0.81667 | 0.82124 | 0.76333 | 0.73333 | 0.915 | 0.76333 | 0.68833 | 0.76667 | 1.0617 | 0.73667 | 0.78833 | 0.76667 | 0.71833 | 0.74 | 0.74833 | 0.85181 | 1 | 0.95 | 0.8 |  | 1.25 | 0.91608 | 1.1183 | 0.93333
| Median |  | 0.8 | 0.81912 | 0.735 | 0.73 | 0.915 | 0.76 | 0.69 | 0.77 | 1.045 | 0.745 | 0.785 | 0.78 | 0.72 | 0.74 | 0.755 | 0.85233 | 1 | 0.95 | 0.8 |  | 1.25 | 0.90532 | 1.13 | 0.935
| Std Dev. |  | 0.040825 | 0.014343 | 0.16597 | 0.024221 | 0.08167 | 0.037771 | 0.01472 | 0.024221 | 0.072503 | 0.030768 | 0.019408 | 0.020656 | 0.017224 | 0.010954 | 0.02137 | 0.017892 | 0 | 0.054772 | 1.2162e-16 |  | 0.054772 | 0.046209 | 0.035449 | 0.008165
| Rel.Std.Dev. |  | 4.999% | 1.746% | 21.74% | 3.303% | 8.926% | 4.948% | 2.138% | 3.159% | 6.829% | 4.177% | 2.462% | 2.694% | 2.398% | 1.48% | 2.856% | 2.1% | 0% | 5.766% | 1.52e-14% |  | 4.382% | 5.044% | 3.17% | 0.8748%
| PDM3 |  | 3.241% | 3.819% | -3.501% | -7.294% | 15.67% | -3.501% | -12.98% | -3.08% | 34.21% | -6.872% | -0.3408% | -3.08% | -9.19% | -6.451% | -5.397% | 7.684% | 26.42% | 20.1% | 1.134% |  | 58.02% | 15.81% | 41.38% | 17.99%
| Z-Score (Absolute) |  | NA | 0.97 | 0.03 | 0.55 | 2.58 | 0.03 | 1.32 | 0.03 | 5.11 | 0.49 | 0.4 | 0.03 | 0.8 | 0.43 | 0.29 | 1.49 | 0.26 | NA | NA | 4.56 | NA | 2.6 | 6.08 | 2.9
Results from laboratories 1, 19, 20 and 25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 471 (Certified Value 0.95 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 25 | 26 | 27 | 29 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 0.9 | 0.88502 | 0.9 | 0.937 | 0.8 | 1.02 | 0.93 | 0.99 | 0.91 | 0.97 | 0.97 | 0.94 | 0.84 | 0.85 | 0.9 | 0.92747 | 0.98 | 0.9 | 0.9 | 0.876 | 1.3228 | 1.0248 | 0.9 | 1.0088 | 0.83 | 1.12 | 1
| 1 | 2 | 0.9 | 0.91675 | 1 | 0.93 | 0.88 | 1.02 | 0.95 | 0.93 | 0.96 | 0.96 | 0.95 | 1 | 0.83 | 0.88 | 0.94 | 0.93185 | 0.98 | 0.9 | 1 | 0.952 | 1.2118 | 1.051 | 1 | 1.0329 | 0.83 | 1.14 | #N/A
| 1 | 3 | 0.9 | 0.91964 | 0.9 | 0.973 | 0.86 | 0.97 | 0.88 | 0.98 | 0.97 | 0.97 | 0.96 | 0.99 | 0.81 | 0.88 | 0.87 | 0.96223 | 0.96 | 0.8 | 1 | 0.986 | 1.3709 | 1.036 | 1 | 1.0281 | 0.84 | 1.1 | 16
| 1 | 4 | 1 | 0.93115 | 1 | 0.956 | 0.77 | 1.05 | 0.92 | 0.97 | 1.01 | 0.97 | 0.98 | 1.02 | 0.82 | 0.83 | 0.87 | 0.91425 | 0.99 | 0.9 | 1 | 1.007 | 1.3562 | 1.0303 | 1 | 0.97662 | 0.8 | 1.12 | 0.94623
| 1 | 5 | 1 | 0.90108 | 0.9 | 0.979 | 0.84 | 0.96 | 0.91 | 0.99 | 1.03 | 0.97 | 0.96 | 0.97 | 0.83 | 0.85 | 0.83 | 0.92872 | 0.98 | 0.9 | 0.9 | 1.032 | 1.3776 | 0.94486 | 1 | 0.99896 | 0.81 | 1.11 | 137
| 1 | 6 | 0.9 | 0.91227 | 0.9 | 0.964 | 0.83 | 1.03 | 0.95 | 1.01 | 1.03 | 1.01 | 0.99 | 0.99 | 0.81 | 0.84 | 0.9 | 0.93234 | 0.98 | 0.8 | 1 | 0.871 | 1.2627 | 1.0901 | 1 | 0.99893 | 0.84 | 1.09
| Mean |  | 0.93333 | 0.91099 | 0.93333 | 0.9565 | 0.83 | 1.0083 | 0.92333 | 0.97833 | 0.985 | 0.975 | 0.96833 | 0.985 | 0.82333 | 0.855 | 0.885 | 0.93281 | 0.97833 | 0.86667 | 0.96667 | 0.954 | 1.317 | 1.0295 | 0.98333 | 1.0074 | 0.825 | 1.1133
| Median |  | 0.9 | 0.91451 | 0.9 | 0.96 | 0.835 | 1.02 | 0.925 | 0.985 | 0.99 | 0.97 | 0.965 | 0.99 | 0.825 | 0.85 | 0.885 | 0.93029 | 0.98 | 0.9 | 1 | 0.969 | 1.3395 | 1.0332 | 1 | 1.0039 | 0.83 | 1.115
| Std Dev. |  | 0.05164 | 0.016055 | 0.05164 | 0.019583 | 0.04 | 0.035449 | 0.026583 | 0.027142 | 0.047223 | 0.017607 | 0.01472 | 0.027386 | 0.012111 | 0.020736 | 0.037283 | 0.015852 | 0.0098319 | 0.05164 | 0.05164 | 0.067667 | 0.066556 | 0.04766 | 0.040825 | 0.020831 | 0.016432 | 0.017512
| Rel.Std.Dev. |  | 5.533% | 1.762% | 5.533% | 2.047% | 4.819% | 3.516% | 2.879% | 2.774% | 4.794% | 1.806% | 1.52% | 2.78% | 1.471% | 2.425% | 4.213% | 1.699% | 1.005% | 5.958% | 5.342% | 7.093% | 5.054% | 4.629% | 4.152% | 2.068% | 1.992% | 1.573%
| PDM3 |  | -1.363% | -3.724% | -1.363% | 1.086% | -12.28% | 6.564% | -2.419% | 3.393% | 4.098% | 3.041% | 2.336% | 4.098% | -12.99% | -9.641% | -6.471% | -1.418% | 3.393% | -8.408% | 2.16% | 0.8216% | 39.18% | 8.802% | 3.922% | 6.464% | -12.81% | 17.66%
| Z-Score (Absolute) |  | NA | 0.97 | NA | 0.2 | 2.33 | 0.67 | 0.76 | 0.17 | 0.28 | 0.11 | 0 | 0.28 | 2.44 | 1.91 | 1.4 | 0.6 | 0.17 | NA | NA | 0.24 | 5.88 | 1.03 | NA | 0.66 | 2.42 | 2.44
Results from laboratories 1, 3, 19, 20 and 25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 471 (Indicative Value 0.085 ppm) | [27 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 27 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.1 | < 0.2 | 0.05 | 0.02 | 0.95 | 0.33 | < 0.1 | < 0.2 | 0.33 | < 0.2 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | 0.02584 | < 0.2 | 0.08 | 0.06 | 0.167 | 0.09 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 0.2 | 0.04 | 0.03 | 0.53 | 0.31 | < 0.1 | < 0.2 | 0.19 | < 0.2 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.02629 | < 0.2 | < 0.05 | < 0.05 | 0.186 | 0.07 | [5 empty] | 2
| 1 | 3 | 0.2 | < 0.2 | < 0.03 | 0.02 | 0.34 | 0.32 | < 0.1 | < 0.2 | 0.13 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.02442 | < 0.2 | 0.06 | 0.07 | 0.257 | 0.08 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 0.2 | 0.04 | 0.01 | 0.19 | 0.33 | < 0.1 | < 0.2 | 0.11 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.02954 | < 0.2 | < 0.05 | < 0.05 | 0.223 | 0.07 | [5 empty] | 0.085053
| 1 | 5 | < 0.1 | < 0.2 | < 0.03 | 0.05 | < 0.05 | 0.32 | < 0.1 | < 0.2 | 0.13 | < 0.2 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.02957 | < 0.2 | < 0.05 | 0.07 | 0.181 | 0.07 | [5 empty] | 31
| 1 | 6 | < 0.1 | < 0.2 | 0.04 | 0.02 | 0.17 | 0.33 | < 0.1 | < 0.2 | 0.15 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.01784 | < 0.2 | < 0.05 | 0.06 | 0.139 | 0.06
| Mean |  | 0.2 |  | 0.0425 | 0.025 | 0.436 | 0.32333 |  |  | 0.17333 |  |  | 0.05 | 0.055 |  |  | 0.025583 |  | 0.07 | 0.065 | 0.19217 | 0.073333
| Median |  | 0.2 |  | 0.04 | 0.02 | 0.34 | 0.325 |  |  | 0.14 |  |  | 0.05 | 0.055 |  |  | 0.026065 |  | 0.07 | 0.065 | 0.1835 | 0.07
| Std Dev. |  |  |  | 0.005 | 0.013784 | 0.32153 | 0.008165 |  |  | 0.081404 |  |  |  | 0.0070711 |  |  | 0.0043227 |  | 0.014142 | 0.0057735 | 0.041897 | 0.010328
| Rel.Std.Dev. |  |  |  | 11.76% | 55.14% | 73.74% | 2.525% |  |  | 46.96% |  |  |  | 12.86% |  |  | 16.9% |  | 20.2% | 8.882% | 21.8% | 14.08%
| PDM3 |  | 135.1% |  | -50.03% | -70.61% | 412.6% | 280.2% |  |  | 103.8% |  |  | -41.21% | -35.33% |  |  | -69.92% |  | -17.7% | -23.58% | 125.9% | -13.78%
| Z-Score (Absolute) |  | 0.59 | 1.22 | 0.46 | 0.67 | 7.99 | 6.87 | 0.04 | 1.22 | 3.08 | 1.22 | 0.67 | 0.57 | 0.42 | 0.67 | 0.67 | 0.66 | 1.22 | 0.3 | 0 | 3.55 | 0.55
| Analytical results for Th in OREAS 471 (Certified Value 23.1 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 26 | 27 | 29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 24 | 21.3357 | 27.48 | 23.4 | 19 | 24.38 | 21.4 | 23.7 | 21.31 | 24.07 | 23.9 | 24.7 | 24.9 | 22.3 | 22.2 | 22.4755 | 22.75 | 16.4 | 24.3 | 21.665 | 23.5392 | 22.9839 | 22.34 | 20.91 |  |  | 1
| 1 | 2 | 23.6 | 21.7243 | 27.12 | 23.1 | 20.2 | 24.21 | 21.2 | 21.99 | 22.2 | 23.37 | 23.4 | 26.4 | 24.1 | 23.2 | 23.7 | 22.4502 | 22.62 | 16.7 | 24.4 | 22.664 | 23.583 | 23.861 | 21.67 | 21.12 |  |  | 15
| 1 | 3 | 23.9 | 21.4658 | 25.29 | 23.9 | 21.4 | 23.28 | 21 | 23 | 22.44 | 23.73 | 23.8 | 26.2 | 23.6 | 23.5 | 23 | 23.2246 | 22.67 | 14 | 24 | 23.364 | 23.3967 | 23.3893 | 22.85 | 20.65 |  |  | 16
| 1 | 4 | 23.6 | 21.9816 | 27.04 | 23.7 | 18.5 | 25.17 | 21.2 | 22.88 | 23.45 | 23.56 | 23.7 | 26.4 | 24.5 | 22.4 | 22.1 | 22.2265 | 23 | 15.9 | 24.5 | 23.773 | 23.5086 | 22.4177 | 23.45 | 20.9 |  |  | 23.1312
| 1 | 5 | 23.9 | 21.6857 | 24.73 | 23.8 | 21.5 | 25.01 | 21.1 | 23.94 | 24.25 | 23.6 | 23 | 25.9 | 24.2 | 23.3 | 20.6 | 22.0708 | 22.92 | 15.6 | 23.6 | 24.593 | 23.1854 | 23.2199 | 22.76 | 20.75 |  |  | 138
| 1 | 6 | 23.7 | 22.108 | 25.5 | 23.4 | 20.2 | 24.25 | 21.2 | 23.52 | 26.47 | 23.78 | 23.7 | 25.7 | 24 | 23 | 22.4 | 21.7149 | 22.86 | 15.1 | 24.1 | 21.38 | 22.9903 | 23.292 | 21.35 | 20.33
| Mean |  | 23.7833 | 21.7168 | 26.1933 | 23.55 | 20.1333 | 24.3833 | 21.1833 | 23.1717 | 23.3533 | 23.685 | 23.5833 | 25.8833 | 24.2167 | 22.95 | 22.3333 | 22.3604 | 22.8033 | 15.6167 | 24.15 | 22.9065 | 23.3672 | 23.194 | 22.4033 | 20.7767
| Median |  | 23.8 | 21.705 | 26.27 | 23.55 | 20.2 | 24.315 | 21.2 | 23.26 | 22.945 | 23.665 | 23.7 | 26.05 | 24.15 | 23.1 | 22.3 | 22.3383 | 22.805 | 15.75 | 24.2 | 23.014 | 23.4526 | 23.256 | 22.55 | 20.825
| Std Dev. |  | 0.17224 | 0.29408 | 1.1549 | 0.30166 | 1.2193 | 0.67479 | 0.13292 | 0.70734 | 1.8367 | 0.23721 | 0.33116 | 0.64317 | 0.4446 | 0.49295 | 1.0386 | 0.50726 | 0.14814 | 0.97451 | 0.32711 | 1.2438 | 0.23329 | 0.4775 | 0.78403 | 0.27097
| Rel.Std.Dev. |  | 0.7242% | 1.354% | 4.409% | 1.281% | 6.056% | 2.767% | 0.6275% | 3.053% | 7.865% | 1.002% | 1.404% | 2.485% | 1.836% | 2.148% | 4.65% | 2.269% | 0.6497% | 6.24% | 1.354% | 5.43% | 0.9984% | 2.059% | 3.5% | 1.304%
| PDM3 |  | 2.819% | -6.114% | 13.24% | 1.811% | -12.96% | 5.413% | -8.421% | 0.175% | 0.9604% | 2.394% | 1.955% | 11.9% | 4.693% | -0.7833% | -3.449% | -3.332% | -1.417% | -32.49% | 4.405% | -0.9714% | 1.02% | 0.2715% | -3.147% | -10.18%
| Z-Score (Absolute) |  | 0.51 | 1.23 | 2.53 | 0.31 | 2.57 | 1.01 | 1.68 | 0.01 | 0.14 | 0.42 | 0.34 | 2.27 | 0.87 | 0.2 | 0.72 | 0.69 | 0.32 | 6.37 | 0.81 | 0.23 | 0.16 | 0.01 | 0.66 | 2.03
| Analytical results for Ti in OREAS 471 (Certified Value 0.531 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 3
3
| 1 | 1 | 0.57 | 0.5156 | 0.437 | 0.49 | 0.5 | 0.55 | 0.501 | 0.5335 | 0.537 | 0.5489 | 0.536 | 0.504 | 0.531 | 0.517 | 0.533 | 0.54079 | 0.531 | 0.526 | 0.566 | 0.5369 | 0.5898 | 0.46 | 8.9216e-05 | 0.535 |  |  | 1
| 1 | 2 | 0.535 | 0.51275 | 0.438 | 0.5 | 0.54 | 0.56 | 0.509 | 0.6174 | 0.547 | 0.5061 | 0.521 | 0.517 | 0.522 | 0.528 | 0.537 | 0.54762 | 0.5406 | 0.541 | 0.567 | 0.5421 | 0.58954 | 0.46 | 9.1151e-05 | 0.542 |  |  | #N/A
| 1 | 3 | 0.559 | 0.51727 | 0.4256 | 0.49 | 0.5 | 0.56 | 0.52 | 0.5248 | 0.549 | 0.5333 | 0.517 | 0.519 | 0.518 | 0.532 | 0.521 | 0.56156 | 0.5199 | 0.534 | 0.561 | 0.5395 | 0.5922 | 0.46 | 9.2413e-05 | 0.536 |  |  | 16
| 1 | 4 | 0.568 | 0.51391 | 0.4391 | 0.5 | 0.5 | 0.56 | 0.508 | 0.6492 | 0.563 | 0.5366 | 0.538 | 0.53 | 0.527 | 0.51 | 0.529 | 0.53705 | 0.5418 | 0.542 | 0.575 | 0.5382 | 0.58446 | 0.46 | 9.3426e-05 | 0.537 |  |  | 0.53068
| 1 | 5 | 0.538 | 0.52341 | 0.4288 | 0.47 | 0.53 | 0.54 | 0.515 | 0.5299 | 0.557 | 0.5089 | 0.521 | 0.52 | 0.519 | 0.523 | 0.512 | 0.53581 | 0.5217 | 0.531 | 0.565 | 0.5434 | 0.58803 | 0.46 | 9.0522e-05 | 0.543 |  |  | 139
| 1 | 6 | 0.58 | 0.53101 | 0.4383 | 0.49 | 0.52 | 0.54 | 0.5 | 0.5419 | 0.528 | 0.5179 | 0.54 | 0.519 | 0.521 | 0.512 | 0.524 | 0.5292 | 0.5354 | 0.531 | 0.563 | 0.5226 | 0.58587 | 0.46 | 9.1595e-05 | 0.539
| Mean |  | 0.55833 | 0.51899 | 0.43447 | 0.49 | 0.515 | 0.55167 | 0.50883 | 0.56612 | 0.54683 | 0.52528 | 0.52883 | 0.51817 | 0.523 | 0.52033 | 0.526 | 0.542 | 0.53173 | 0.53417 | 0.56617 | 0.53712 | 0.58832 | 0.46 | 9.1387e-05 | 0.53867
| Median |  | 0.5635 | 0.51643 | 0.4375 | 0.49 | 0.51 | 0.555 | 0.5085 | 0.5377 | 0.548 | 0.5256 | 0.5285 | 0.519 | 0.5215 | 0.52 | 0.5265 | 0.53892 | 0.5332 | 0.5325 | 0.5655 | 0.53885 | 0.58879 | 0.46 | 9.1373e-05 | 0.538
| Std Dev. |  | 0.018206 | 0.0069757 | 0.0057584 | 0.010954 | 0.017607 | 0.0098319 | 0.0077825 | 0.053295 | 0.012813 | 0.016977 | 0.010226 | 0.0083287 | 0.00502 | 0.0088242 | 0.0089889 | 0.011334 | 0.0093234 | 0.0062423 | 0.0048339 | 0.0075093 | 0.0028195 | 0 | 1.4676e-06 | 0.003266
| Rel.Std.Dev. |  | 3.261% | 1.344% | 1.325% | 2.236% | 3.419% | 1.782% | 1.529% | 9.414% | 2.343% | 3.232% | 1.934% | 1.607% | 0.9598% | 1.696% | 1.709% | 2.091% | 1.753% | 1.169% | 0.8538% | 1.398% | 0.4793% | 0% | 1.606% | 0.6063%
| PDM3 |  | 5.21% | -2.203% | -18.13% | -7.666% | -2.955% | 3.954% | -4.117% | 6.677% | 3.043% | -1.017% | -0.3485% | -2.358% | -1.448% | -1.95% | -0.8824% | 2.133% | 0.198% | 0.6565% | 6.687% | 1.212% | 10.86% | -13.32% | -99.98% | 1.505%
| Z-Score (Absolute) |  | 1.51 | 0.5 | 4.83 | 1.99 | 0.71 | 1.17 | 1.02 | 1.91 | 0.92 | 0.18 | 0 | 0.55 | 0.3 | 0.44 | 0.15 | 0.67 | 0.15 | 0.27 | 1.91 | 0.42 | 3.05 | 3.52 | NA | 0.5
Results from laboratory 1.26 were removed due to their systematic bias.
| Analytical results for Tl in OREAS 471 (Certified Value 0.68 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.69 | 0.68269 | 0.669 | 0.7 | 0.69 | 0.62 | 0.7 | 0.71 | 0.68 | 0.71 | 0.65 | 0.72 | 0.67 | 0.67 | 0.66 | 0.71598 | 0.73 | 0.64 | 0.75 | 0.62 | 0.42109 | 0.75 | 0.64266 | 0.58 | 0.57 |  | 1
| 1 | 2 | 0.67 | 0.69507 | 0.655 | 0.7 | 0.73 | 0.61 | 0.7 | 0.71 | 0.71 | 0.71 | 0.66 | 0.77 | 0.64 | 0.69 | 0.69 | 0.73581 | 0.72 | 0.62 | 0.75 | 0.617 | 0.4375 | 0.75 | 0.63399 | 0.58 | 0.57 |  | 16
| 1 | 3 | 0.68 | 0.67485 | 0.625 | 0.71 | 0.69 | 0.58 | 0.69 | 0.73 | 0.73 | 0.73 | 0.64 | 0.76 | 0.64 | 0.7 | 0.64 | 0.73445 | 0.73 | 0.67 | 0.74 | 0.644 | 0.41569 | 0.74 | 0.64982 | 0.59 | 0.57 |  | 16
| 1 | 4 | 0.68 | 0.69372 | 0.667 | 0.7 | 0.66 | 0.63 | 0.69 | 0.69 | 0.73 | 0.73 | 0.66 | 0.77 | 0.67 | 0.67 | 0.66 | 0.71552 | 0.71 | 0.69 | 0.74 | 0.635 | 0.4362 | 0.74 | 0.62829 | 0.56 | 0.57 |  | 0.67882
| 1 | 5 | 0.68 | 0.68116 | 0.61 | 0.73 | 0.7 | 0.64 | 0.68 | 0.73 | 0.76 | 0.71 | 0.64 | 0.72 | 0.67 | 0.69 | 0.65 | 0.69814 | 0.72 | 0.65 | 0.73 | 0.659 | 0.44527 | 0.75 | 0.65402 | 0.58 | 0.57 |  | 140
| 1 | 6 | 0.69 | 0.68084 | 0.608 | 0.71 | 0.68 | 0.6 | 0.71 | 0.73 | 0.76 | 0.72 | 0.69 | 0.75 | 0.63 | 0.69 | 0.68 | 0.7189 | 0.75 | 0.66 | 0.75 | 0.58 | 0.47499 | 0.74 | 0.64549 | 0.59 | 0.57
| Mean |  | 0.68167 | 0.68472 | 0.639 | 0.70833 | 0.69167 | 0.61333 | 0.695 | 0.71667 | 0.72833 | 0.71833 | 0.65667 | 0.74833 | 0.65333 | 0.685 | 0.66333 | 0.7198 | 0.72667 | 0.655 | 0.74333 | 0.62583 | 0.43846 | 0.745 | 0.64238 | 0.58 | 0.57
| Median |  | 0.68 | 0.68192 | 0.64 | 0.705 | 0.69 | 0.615 | 0.695 | 0.72 | 0.73 | 0.715 | 0.655 | 0.755 | 0.655 | 0.69 | 0.66 | 0.71744 | 0.725 | 0.655 | 0.745 | 0.6275 | 0.43685 | 0.745 | 0.64408 | 0.58 | 0.57
| Std Dev. |  | 0.0075277 | 0.0079673 | 0.028064 | 0.01169 | 0.023166 | 0.021602 | 0.010488 | 0.01633 | 0.030605 | 0.0098319 | 0.018619 | 0.023166 | 0.018619 | 0.012247 | 0.018619 | 0.013956 | 0.013663 | 0.02429 | 0.008165 | 0.027301 | 0.021008 | 0.0054772 | 0.0096909 | 0.010954 | 0
| Rel.Std.Dev. |  | 1.104% | 1.164% | 4.392% | 1.65% | 3.349% | 3.522% | 1.509% | 2.279% | 4.202% | 1.369% | 2.835% | 3.096% | 2.85% | 1.788% | 2.807% | 1.939% | 1.88% | 3.708% | 1.098% | 4.362% | 4.791% | 0.7352% | 1.509% | 1.889% | 0%
| PDM3 |  | 0.419% | 0.8689% | -5.866% | 4.347% | 1.892% | -9.647% | 2.383% | 5.575% | 7.294% | 5.821% | -3.264% | 10.24% | -3.755% | 0.9101% | -2.282% | 6.037% | 7.048% | -3.509% | 9.503% | -7.806% | -35.41% | 9.749% | -5.368% | -14.56% | -16.03%
| Z-Score (Absolute) |  | 0.06 | 0 | 0.88 | 0.45 | 0.13 | 1.37 | 0.2 | 0.61 | 0.84 | 0.65 | 0.54 | 1.22 | 0.6 | 0.01 | 0.41 | 0.67 | 0.81 | 0.57 | 1.13 | 1.13 | 4.73 | 1.16 | 0.81 | 2.01 | 2.21
| Analytical results for Tm in OREAS 471 (Certified Value 0.3 ppm) | [27 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.3 | 0.28556 | 0.28 | 0.324 | 0.33314 | 0.27 | 0.3 | 0.32 | 0.29 | 0.34 | 0.33 | 0.3 | 0.26 | 0.27 | 0.28 | 0.33574 | 0.33 | 0.3 | 0.3 | 0.265 | 0.27924 | 0.25719 | 0.30339 | 0.35 | < 1 |  | 1
| 1 | 2 | 0.3 | 0.29173 | 0.28 | 0.318 | 0.36726 | 0.3 | 0.3 | 0.3 | 0.3 | 0.34 | 0.31 | 0.32 | 0.25 | 0.28 | 0.3 | 0.33521 | 0.32 | 0.3 | 0.3 | 0.284 | 0.264 | 0.26447 | 0.31216 | 0.32 | < 1 |  | #N/A
| 1 | 3 | 0.3 | 0.29409 | 0.26 | 0.325 | 0.35699 | 0.28 | 0.3 | 0.31 | 0.31 | 0.33 | 0.31 | 0.32 | 0.25 | 0.28 | 0.29 | 0.34774 | 0.31 | 0.3 | 0.3 | 0.296 | 0.30519 | 0.26673 | 0.3194 | 0.34 | < 1 |  | 16
| 1 | 4 | 0.3 | 0.29171 | 0.28 | 0.318 | 0.31952 | 0.29 | 0.3 | 0.32 | 0.31 | 0.34 | 0.32 | 0.32 | 0.27 | 0.27 | 0.28 | 0.33409 | 0.32 | 0.3 | 0.3 | 0.308 | 0.28656 | 0.23505 | 0.29827 | 0.33 | < 1 |  | 0.30278
| 1 | 5 | 0.3 | 0.28791 | 0.26 | 0.326 | 0.36248 | 0.28 | 0.3 | 0.33 | 0.32 | 0.32 | 0.31 | 0.31 | 0.26 | 0.27 | 0.27 | 0.33029 | 0.32 | 0.3 | 0.3 | 0.317 | 0.30633 | 0.24616 | 0.30685 | 0.33 | < 1 |  | 141
| 1 | 6 | 0.3 | 0.28743 | 0.27 | 0.316 | 0.34111 | 0.28 | 0.3 | 0.36 | 0.33 | 0.32 | 0.32 | 0.31 | 0.27 | 0.27 | 0.29 | 0.33856 | 0.33 | 0.3 | 0.3 | 0.259 | 0.27855 | 0.25424 | 0.30277 | 0.34 | < 1
| Mean |  | 0.3 | 0.28974 | 0.27167 | 0.32117 | 0.34675 | 0.28333 | 0.3 | 0.32333 | 0.31 | 0.33167 | 0.31667 | 0.31333 | 0.26 | 0.27333 | 0.285 | 0.33694 | 0.32167 | 0.3 | 0.3 | 0.28817 | 0.28664 | 0.25397 | 0.30714 | 0.335
| Median |  | 0.3 | 0.28981 | 0.275 | 0.321 | 0.34905 | 0.28 | 0.3 | 0.32 | 0.31 | 0.335 | 0.315 | 0.315 | 0.26 | 0.27 | 0.285 | 0.33547 | 0.32 | 0.3 | 0.3 | 0.29 | 0.2829 | 0.25571 | 0.30512 | 0.335
| Std Dev. |  | 0 | 0.0032508 | 0.0098319 | 0.0043089 | 0.01861 | 0.010328 | 0 | 0.020656 | 0.014142 | 0.0098319 | 0.008165 | 0.008165 | 0.0089443 | 0.005164 | 0.010488 | 0.0059337 | 0.0075277 | 0 | 0 | 0.023198 | 0.01652 | 0.011853 | 0.0075816 | 0.010488
| Rel.Std.Dev. |  | 0% | 1.122% | 3.619% | 1.342% | 5.367% | 3.645% | 0% | 6.388% | 4.562% | 2.964% | 2.578% | 2.606% | 3.44% | 1.889% | 3.68% | 1.761% | 2.34% | 0% | 0% | 8.05% | 5.763% | 4.667% | 2.468% | 3.131%
| PDM3 |  | -0.9167% | -4.307% | -10.27% | 6.074% | 14.52% | -6.421% | -0.9167% | 6.79% | 2.386% | 9.542% | 4.588% | 3.487% | -14.13% | -9.724% | -5.871% | 11.28% | 6.239% | -0.9167% | -0.9167% | -4.825% | -5.329% | -16.12% | 1.442% | 10.64%
| Z-Score (Absolute) |  | NA | 0.58 | 1.11 | 0.32 | 1.06 | 0.77 | NA | 0.38 | 0 | 0.63 | 0.19 | 0.1 | 1.44 | 1.06 | 0.72 | 0.78 | 0.34 | NA | NA | 0.63 | 0.67 | 1.62 | 0.08 | 0.72 | 5.49
Results from laboratories 1, 8, 19 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 471 (Certified Value 2.44 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 26 | 27 | 29 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.4 | 2.3486 | 2.27 | 2.37 | 2.2 | 2.75 | 2.37 | 2.57 | 2.37 | 2.63 | 2.6 | 2.5 | 2.4 | 2.5 | 2.4 | 2.5732 | 2.5 | 2.4 | 2.5 | 2.445 | 1.6476 | 2.3409 | 2.7 | 2.17 |  |  | 1
| 1 | 2 | 2.4 | 2.3864 | 2.18 | 2.38 | 2.4 | 2.65 | 2.37 | 2.5 | 2.38 | 2.47 | 2.5 | 2.6 | 2.3 | 2.6 | 2.5 | 2.5753 | 2.51 | 2.5 | 2.5 | 2.325 | 1.6385 | 2.2716 | 2.6 | 2.19 |  |  | #N/A
| 1 | 3 | 2.4 | 2.3007 | 2.06 | 2.46 | 2.3 | 2.67 | 2.21 | 2.52 | 2.41 | 2.54 | 2.5 | 2.7 | 2.2 | 2.6 | 2.4 | 2.6725 | 2.49 | 2.3 | 2.5 | 2.397 | 1.7204 | 2.4069 | 2.7 | 2.18 |  |  | 16
| 1 | 4 | 2.4 | 2.4205 | 2.22 | 2.38 | 2.2 | 2.79 | 2.36 | 2.52 | 2.54 | 2.57 | 2.6 | 2.7 | 2.3 | 2.5 | 2.3 | 2.5267 | 2.56 | 2.4 | 2.6 | 2.422 | 1.6773 | 2.2134 | 2.6 | 2.15 |  |  | 2.4423
| 1 | 5 | 2.3 | 2.3687 | 2.02 | 2.43 | 2.3 | 2.73 | 2.34 | 2.61 | 2.58 | 2.47 | 2.5 | 2.6 | 2.3 | 2.6 | 2.3 | 2.5366 | 2.55 | 2.4 | 2.5 | 2.467 | 1.6689 | 2.2704 | 2.6 | 2.13 |  |  | 142
| 1 | 6 | 2.4 | 2.5075 | 2.1 | 2.42 | 2.3 | 2.65 | 2.33 | 2.6 | 2.83 | 2.51 | 2.6 | 2.6 | 2.3 | 2.5 | 2.4 | 2.4902 | 2.59 | 2.4 | 2.5 | 2.384 | 1.6873 | 2.4273 | 2.6 | 2.11
| Mean |  | 2.3833 | 2.3887 | 2.1417 | 2.4067 | 2.2833 | 2.7067 | 2.33 | 2.5533 | 2.5183 | 2.5317 | 2.55 | 2.6167 | 2.3 | 2.55 | 2.3833 | 2.5624 | 2.5333 | 2.4 | 2.5167 | 2.4067 | 1.6733 | 2.3217 | 2.6333 | 2.155
| Median |  | 2.4 | 2.3775 | 2.14 | 2.4 | 2.3 | 2.7 | 2.35 | 2.545 | 2.475 | 2.525 | 2.55 | 2.6 | 2.3 | 2.55 | 2.4 | 2.5549 | 2.53 | 2.4 | 2.5 | 2.4095 | 1.6731 | 2.3063 | 2.6 | 2.16
| Std Dev. |  | 0.040825 | 0.070536 | 0.097245 | 0.03559 | 0.075277 | 0.058538 | 0.060992 | 0.046332 | 0.17566 | 0.062102 | 0.054772 | 0.075277 | 0.063246 | 0.054772 | 0.075277 | 0.062517 | 0.039328 | 0.063246 | 0.040825 | 0.050242 | 0.029394 | 0.084434 | 0.05164 | 0.030822
| Rel.Std.Dev. |  | 1.713% | 2.953% | 4.541% | 1.479% | 3.297% | 2.163% | 2.618% | 1.815% | 6.975% | 2.453% | 2.148% | 2.877% | 2.75% | 2.148% | 3.158% | 2.44% | 1.552% | 2.635% | 1.622% | 2.088% | 1.757% | 3.637% | 1.961% | 1.43%
| PDM3 |  | -2.414% | -2.194% | -12.31% | -1.459% | -6.509% | 10.82% | -4.598% | 4.546% | 3.113% | 3.659% | 4.41% | 7.139% | -5.826% | 4.41% | -2.414% | 4.919% | 3.727% | -1.732% | 3.045% | -1.459% | -31.49% | -4.936% | 7.822% | -11.76%
| Z-Score (Absolute) |  | 0.13 | 0.1 | 1.44 | 0 | 0.67 | 1.63 | 0.42 | 0.8 | 0.61 | 0.68 | 0.78 | 1.14 | 0.58 | 0.78 | 0.13 | 0.85 | 0.69 | 0.04 | 0.6 | 0 | 3.98 | 0.46 | 1.23 | 1.37
| Analytical results for V in OREAS 471 (Certified Value 102 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 100 | 95.9287 | 114 | 103 | 94 | 97 | 101 | 106 | 100 | 104 | 104 | 99 | 103 | 98 | 103 | 104.6065 | 106 | 102 | 106 | 103.844 | 103.9921 | 101 | 98.4609 | 105 |  |  | 1
| 1 | 2 | 98 | 96.3455 | 115 | 107 | 100 | 100 | 101 | 104 | 100 | 99 | 101 | 102 | 98 | 102 | 105 | 105.2621 | 106 | 105 | 106 | 111.486 | 106.5124 | 101 | 100.2285 | 106 |  |  | #N/A
| 1 | 3 | 97 | 95.8743 | 114 | 101 | 97 | 99 | 101 | 95 | 102 | 103 | 99 | 104 | 100 | 101 | 100 | 107.6275 | 104 | 103 | 104 | 101.223 | 106.066 | 102 | 100.0198 | 105 |  |  | 16
| 1 | 4 | 101 | 96.4325 | 115 | 104 | 91 | 100 | 100 | 104 | 103 | 99 | 103 | 109 | 100 | 97 | 100 | 104.0195 | 108 | 105 | 107 | 107.712 | 107.7307 | 103 | 99.5927 | 106 |  |  | 101.9604
| 1 | 5 | 99 | 97.3184 | 113 | 103 | 98 | 97 | 101 | 101 | 103 | 103 | 101 | 102 | 100 | 99 | 100 | 102.7435 | 105 | 103 | 105 | 99.995 | 108.3259 | 102 | 98.0803 | 107 |  |  | 143
| 1 | 6 | 99 | 98.8852 | 115 | 98.4 | 97 | 98 | 98 | 105 | 97 | 106 | 105 | 100 | 101 | 100 | 103 | 101.5652 | 107 | 104 | 105 | 99.907 | 105.7126 | 102 | 99.4431 | 106
| Mean |  | 99 | 96.7974 | 114.3333 | 102.7333 | 96.1667 | 98.5 | 100.3333 | 102.5 | 100.8333 | 102.3333 | 102.1667 | 102.6667 | 100.3333 | 99.5 | 101.8333 | 104.3041 | 106 | 103.6667 | 105.5 | 104.0278 | 106.3899 | 101.8333 | 99.3042 | 105.8333
| Median |  | 99 | 96.389 | 114.5 | 103 | 97 | 98.5 | 101 | 104 | 101 | 103 | 102 | 102 | 100 | 99.5 | 101.5 | 104.313 | 106 | 103.5 | 105.5 | 102.5335 | 106.2892 | 102 | 99.5179 | 106
| Std Dev. |  | 1.4142 | 1.1467 | 0.8165 | 2.8891 | 3.1885 | 1.3784 | 1.2111 | 4.0373 | 2.3166 | 2.8048 | 2.2286 | 3.559 | 1.633 | 1.8708 | 2.137 | 2.1011 | 1.4142 | 1.2111 | 1.0488 | 4.6977 | 1.541 | 0.75277 | 0.85766 | 0.75277
| Rel.Std.Dev. |  | 1.428% | 1.185% | 0.7141% | 2.812% | 3.316% | 1.399% | 1.207% | 3.939% | 2.297% | 2.741% | 2.181% | 3.467% | 1.628% | 1.88% | 2.099% | 2.014% | 1.334% | 1.168% | 0.9941% | 4.516% | 1.448% | 0.7392% | 0.8637% | 0.7113%
| PDM3 |  | -2.904% | -5.064% | 12.13% | 0.758% | -5.682% | -3.394% | -1.596% | 0.5292% | -1.105% | 0.3657% | 0.2023% | 0.6926% | -1.596% | -2.413% | -0.1247% | 2.299% | 3.962% | 1.673% | 3.472% | 2.028% | 4.344% | -0.1247% | -2.605% | 3.798%
| Z-Score (Absolute) |  | 1.1 | 1.85 | 4.1 | 0.16 | 2.07 | 1.27 | 0.65 | 0.08 | 0.48 | 0.03 | 0.03 | 0.14 | 0.65 | 0.93 | 0.14 | 0.7 | 1.27 | 0.48 | 1.1 | 0.6 | 1.41 | 0.14 | 1 | 1.22
| Analytical results for W in OREAS 471 (Certified Value 3.99 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 3.7 | 3.7272 | 4.04 | 4.2 | 3.8 | 2.66 | 4.6 | 4 | 4 | 4.2 | 4 | 3.8 | 3.8 | 3.6 | 3.8 | 4.0544 | 4.1 | 3.8 | 4.2 | 4.187 | 3.8447 | 4.1 | 4.3701 | 4.1 |  |  | 1
| 1 | 2 | 3.8 | 3.86 | 3.94 | 4.1 | 3.9 | 2.67 | 4.4 | 4 | 3.8 | 3.8 | 3.9 | 4.1 | 3.7 | 3.7 | 3.9 | 4.111 | 4.1 | 4.1 | 4.2 | 3.945 | 3.8124 | 3.9 | 4.3707 | 4.1 |  |  | #N/A
| 1 | 3 | 3.8 | 3.9241 | 3.76 | 4.2 | 3.7 | 2.53 | 4.4 | 4 | 4 | 4.1 | 3.9 | 4.1 | 3.6 | 3.7 | 3.7 | 4.1539 | 4.2 | 4.1 | 4.2 | 4.251 | 3.8482 | 3.9 | 4.6391 | 4 |  |  | 16
| 1 | 4 | 3.8 | 3.8952 | 4.05 | 4.1 | 3.6 | 2.76 | 4.5 | 3.9 | 4 | 4.2 | 3.9 | 4.1 | 3.7 | 3.5 | 3.8 | 4.0636 | 4.1 | 4.2 | 4.3 | 3.938 | 3.744 | 4 | 4.1711 | 4.1 |  |  | 3.9908
| 1 | 5 | 3.8 | 3.8264 | 3.68 | 4.1 | 4.1 | 2.71 | 4.2 | 4.1 | 4.3 | 4.2 | 3.8 | 3.9 | 3.7 | 3.6 | 3.5 | 4.0052 | 4.1 | 4 | 4.1 | 4.038 | 3.7431 | 4 | 4.1896 | 4 |  |  | 144
| 1 | 6 | 3.7 | 3.9288 | 3.77 | 4.1 | 3.8 | 2.79 | 4.8 | 4.1 | 4.5 | 3.9 | 3.9 | 4 | 3.7 | 3.6 | 3.8 | 4.0999 | 4.2 | 4 | 4.2 | 4.076 | 3.7673 | 4 | 4.4879 | 4
| Mean |  | 3.7667 | 3.8603 | 3.8733 | 4.1333 | 3.8167 | 2.6867 | 4.4833 | 4.0167 | 4.1 | 4.0667 | 3.9 | 4 | 3.7 | 3.6167 | 3.75 | 4.0813 | 4.1333 | 4.0333 | 4.2 | 4.0725 | 3.7933 | 3.9833 | 4.3714 | 4.05
| Median |  | 3.8 | 3.8776 | 3.855 | 4.1 | 3.8 | 2.69 | 4.45 | 4 | 4 | 4.15 | 3.9 | 4.05 | 3.7 | 3.6 | 3.8 | 4.0817 | 4.1 | 4.05 | 4.2 | 4.057 | 3.7899 | 4 | 4.3704 | 4.05
| Std Dev. |  | 0.05164 | 0.075952 | 0.15769 | 0.05164 | 0.17224 | 0.091797 | 0.20412 | 0.075277 | 0.25298 | 0.17512 | 0.063246 | 0.12649 | 0.063246 | 0.075277 | 0.13784 | 0.051644 | 0.05164 | 0.13663 | 0.063246 | 0.1269 | 0.048242 | 0.075277 | 0.17799 | 0.054772
| Rel.Std.Dev. |  | 1.371% | 1.968% | 4.071% | 1.249% | 4.513% | 3.417% | 4.553% | 1.874% | 6.17% | 4.306% | 1.622% | 3.162% | 1.709% | 2.081% | 3.676% | 1.265% | 1.249% | 3.387% | 1.506% | 3.116% | 1.272% | 1.89% | 4.072% | 1.352%
| PDM3 |  | -5.617% | -3.271% | -2.944% | 3.571% | -4.364% | -32.68% | 12.34% | 0.6477% | 2.736% | 1.901% | -2.276% | 0.2301% | -7.287% | -9.375% | -6.034% | 2.268% | 3.571% | 1.065% | 5.242% | 2.047% | -4.95% | -0.1875% | 9.537% | 1.483%
| Z-Score (Absolute) |  | 1.25 | 0.77 | 0.7 | 0.65 | 0.99 | 6.86 | 2.46 | 0.04 | 0.48 | 0.3 | 0.56 | 0.04 | 1.6 | 2.03 | 1.34 | 0.38 | 0.65 | 0.13 | 0.99 | 0.33 | 1.12 | 0.13 | 1.88 | 0.22
| Analytical results for Y in OREAS 471 (Certified Value 22.6 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 19.5 | 20.3043 | 23.06 | 21.6 | 17.7 | 22.8 | 21.4 | 22.57 | 21.7 | 23.05 | 23.8 | 23 | 24.3 | 23.2 | 22.1 | 25.2959 | 22.44 | 20 | 23 | 20.242 | 28.5516 | 27.3579 | 20.9074 | 25.4 | 21.057 |  | 1
| 1 | 2 | 19.6 | 20.4675 | 23.25 | 21.7 | 19.7 | 22.8 | 21.2 | 20.5 | 22.7 | 22.28 | 23.6 | 24.6 | 23.2 | 23.8 | 22.8 | 25.1862 | 23.06 | 19.6 | 23.4 | 22.974 | 28.7191 | 27.313 | 21.4551 | 25.3 | 21.561 |  | 19
| 1 | 3 | 20.1 | 20.566 | 22.78 | 20.8 | 20.4 | 22.8 | 21.7 | 22.22 | 22.1 | 22.84 | 23.9 | 24.9 | 22.5 | 23.9 | 22.5 | 25.6058 | 22.48 | 17.1 | 23.3 | 22.914 | 27.8557 | 27.2492 | 21.5624 | 25 | 21.498 |  | 16
| 1 | 4 | 19.9 | 20.5601 | 23.09 | 21.2 | 17.7 | 23.6 | 22.1 | 22.41 | 23.9 | 23.08 | 24.1 | 24.4 | 23.7 | 22.9 | 22.4 | 25.0896 | 22.76 | 19.1 | 23.3 | 23.417 | 27.868 | 26.8043 | 20.5857 | 25.4 | 21.033 |  | 22.5706
| 1 | 5 | 19.5 | 20.9485 | 22.83 | 21.4 | 20 | 22.8 | 21.6 | 22.61 | 24.1 | 22.97 | 23.2 | 23.7 | 23.9 | 23.9 | 20.4 | 25.209 | 22.57 | 18.7 | 23.1 | 24.049 | 28.7073 | 27.113 | 21.1215 | 25.9 | 21.435 |  | 145
| 1 | 6 | 20.3 | 21.201 | 23.1 | 20.5 | 18.7 | 22.8 | 21.7 | 22.95 | 25 | 23.12 | 24.4 | 24.4 | 23.1 | 23.9 | 22.6 | 25.7721 | 22.76 | 17.8 | 23.2 | 19.621 | 29.2025 | 27.0544 | 21.2777 | 25.5 | 20.947
| Mean |  | 19.8167 | 20.6746 | 23.0183 | 21.2 | 19.0333 | 22.9333 | 21.6167 | 22.21 | 23.25 | 22.89 | 23.8333 | 24.1667 | 23.45 | 23.6 | 22.1333 | 25.3598 | 22.6783 | 18.7167 | 23.2167 | 22.2028 | 28.484 | 27.1486 | 21.1516 | 25.4167 | 21.2552
| Median |  | 19.75 | 20.5631 | 23.075 | 21.3 | 19.2 | 22.8 | 21.65 | 22.49 | 23.3 | 23.01 | 23.85 | 24.4 | 23.45 | 23.85 | 22.45 | 25.2525 | 22.665 | 18.9 | 23.25 | 22.944 | 28.6295 | 27.1811 | 21.1996 | 25.4 | 21.246
| Std Dev. |  | 0.33714 | 0.33379 | 0.1786 | 0.46904 | 1.1759 | 0.3266 | 0.30605 | 0.87194 | 1.2834 | 0.31483 | 0.41312 | 0.69474 | 0.6442 | 0.43818 | 0.88015 | 0.26853 | 0.23103 | 1.098 | 0.1472 | 1.8162 | 0.52921 | 0.20476 | 0.36265 | 0.29269 | 0.27145
| Rel.Std.Dev. |  | 1.701% | 1.614% | 0.7759% | 2.212% | 6.178% | 1.424% | 1.416% | 3.926% | 5.52% | 1.375% | 1.733% | 2.875% | 2.747% | 1.857% | 3.977% | 1.059% | 1.019% | 5.867% | 0.634% | 8.18% | 1.858% | 0.7542% | 1.715% | 1.152% | 1.277%
| PDM3 |  | -12.2% | -8.401% | 1.984% | -6.073% | -15.67% | 1.607% | -4.227% | -1.598% | 3.01% | 1.415% | 5.594% | 7.071% | 3.896% | 4.561% | -1.938% | 12.36% | 0.4771% | -17.08% | 2.862% | -1.63% | 26.2% | 20.28% | -6.287% | 12.61% | -5.828%
| Z-Score (Absolute) |  | 1.63 | 1.17 | 0.07 | 0.89 | 2.04 | 0.04 | 0.67 | 0.36 | 0.19 | 0 | 0.5 | 0.68 | 0.3 | 0.38 | 0.4 | 1.31 | 0.11 | 2.21 | 0.17 | 0.36 | 2.96 | 2.26 | 0.92 | 1.34 | 0.87
| Analytical results for Yb in OREAS 471 (Certified Value 1.96 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 24 | 26 | 27 | 29 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 2 | 1.8562 | 1.86 | 1.93 | 1.7 | 2 | 1.8 | 2.11 | 1.9 | 1.98 | 2.14 | 1.9 | 2 | 1.83 | 1.82 | 2.153 | 1.99 | 1.9 | 2.1 | 1.868 | 1.8532 | 1.9769 | 2.0138 | 2.07 | 2.07 |  | 1
| 1 | 2 | 2 | 1.8877 | 1.82 | 1.87 | 1.9 | 1.96 | 1.9 | 2.05 | 2 | 2.04 | 1.97 | 1.99 | 1.92 | 1.88 | 1.89 | 2.2572 | 2 | 1.8 | 2.1 | 1.976 | 1.7587 | 1.8875 | 2.0441 | 2.09 | 2.09 |  | #N/A
| 1 | 3 | 2.1 | 1.8975 | 1.72 | 1.93 | 1.8 | 1.87 | 1.9 | 2.09 | 2 | 2.02 | 1.96 | 2.02 | 1.84 | 1.86 | 1.78 | 2.2591 | 1.89 | 1.7 | 2.1 | 2.045 | 2.0246 | 1.9435 | 2.1121 | 2.03 | 2.08 |  | 16
| 1 | 4 | 2.2 | 1.8839 | 1.8 | 1.9 | 1.6 | 1.95 | 1.8 | 2.06 | 2.1 | 2.1 | 2.05 | 2.01 | 1.98 | 1.78 | 1.8 | 2.2036 | 1.99 | 1.8 | 2.1 | 2.095 | 1.8766 | 1.9912 | 1.9637 | 2.09 | 2.01 |  | 1.9605
| 1 | 5 | 2 | 1.9169 | 1.68 | 1.96 | 1.8 | 2.02 | 1.7 | 2.14 | 2.2 | 2.06 | 1.98 | 2 | 1.98 | 1.82 | 1.7 | 2.1549 | 1.97 | 1.7 | 2.1 | 2.174 | 2.034 | 1.9455 | 1.9774 | 2.03 | 2.06 |  | 146
| 1 | 6 | 2 | 1.9345 | 1.71 | 1.89 | 1.7 | 2 | 1.9 | 2.13 | 2.4 | 2.15 | 2.09 | 2 | 1.92 | 1.8 | 1.84 | 2.156 | 2.06 | 1.7 | 2.1 | 1.867 | 1.8218 | 2.0433 | 1.9988 | 2.04 | 2.07
| Mean |  | 2.05 | 1.8961 | 1.765 | 1.9133 | 1.75 | 1.9667 | 1.8333 | 2.0967 | 2.1 | 2.0583 | 2.0317 | 1.9867 | 1.94 | 1.8283 | 1.805 | 2.1973 | 1.9833 | 1.7667 | 2.1 | 2.0042 | 1.8948 | 1.9647 | 2.0183 | 2.0583 | 2.0633
| Median |  | 2 | 1.8926 | 1.76 | 1.915 | 1.75 | 1.98 | 1.85 | 2.1 | 2.05 | 2.05 | 2.015 | 2 | 1.95 | 1.825 | 1.81 | 2.1798 | 1.99 | 1.75 | 2.1 | 2.0105 | 1.8649 | 1.9612 | 2.0063 | 2.055 | 2.07
| Std Dev. |  | 0.083666 | 0.027274 | 0.071484 | 0.03266 | 0.10488 | 0.054283 | 0.08165 | 0.036697 | 0.17889 | 0.060139 | 0.073598 | 0.043665 | 0.05933 | 0.037103 | 0.063797 | 0.050816 | 0.055015 | 0.08165 | 0 | 0.12402 | 0.11149 | 0.052548 | 0.053893 | 0.028577 | 0.028048
| Rel.Std.Dev. |  | 4.081% | 1.438% | 4.05% | 1.707% | 5.993% | 2.76% | 4.454% | 1.75% | 8.518% | 2.922% | 3.623% | 2.198% | 3.058% | 2.029% | 3.534% | 2.313% | 2.774% | 4.622% | 0% | 6.188% | 5.884% | 2.675% | 2.67% | 1.388% | 1.359%
| PDM3 |  | 4.566% | -3.283% | -9.971% | -2.405% | -10.74% | 0.3154% | -6.486% | 6.946% | 7.116% | 4.991% | 3.631% | 1.336% | -1.045% | -6.741% | -7.931% | 12.08% | 1.166% | -9.886% | 7.116% | 2.228% | -3.348% | 0.2128% | 2.951% | 4.991% | 5.246%
| Z-Score (Absolute) |  | 0.56 | 0.74 | 1.84 | 0.59 | 1.97 | 0.14 | 1.26 | 0.96 | 0.98 | 0.63 | 0.41 | 0.03 | 0.37 | 1.31 | 1.5 | 1.8 | 0 | 1.83 | 0.98 | 0.18 | 0.75 | 0.16 | 0.3 | 0.63 | 0.67
| Analytical results for Zn in OREAS 471 (Certified Value 77 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 0
1
| 1 | 1 | 76 | 73.4269 | 78 | 79.4 | 69 | 74 | 76 | 77 | 71 | 76 | 80 | 72 | 79 | 75 | 77 | 81.6216 | 77 | 78.9 | 73.6 | 74.426 | 81.27 | 78.5 | 70.8091 | 76 |  |  | 1
| 1 | 2 | 72 | 73.3064 | 78 | 78.2 | 75 | 78 | 73 | 77 | 74 | 76 | 79 | 74 | 75 | 77 | 78 | 76.311 | 76 | 79.4 | 75.7 | 84.798 | 82.215 | 78 | 72.9943 | 77 |  |  | #N/A
| 1 | 3 | 75 | 73.6377 | 77 | 79 | 70 | 76 | 74 | 74 | 74 | 75 | 76 | 76 | 77 | 77 | 74 | 77.8933 | 76 | 78.2 | 74.8 | 76.568 | 80.913 | 78.5 | 74.3463 | 76 |  |  | 16
| 1 | 4 | 77 | 73.3013 | 77 | 79.9 | 69 | 79 | 76 | 78 | 76 | 80 | 80 | 80 | 77 | 74 | 75 | 77.4425 | 77 | 78.1 | 74.7 | 82.571 | 83.454 | 78.9 | 76.6065 | 77 |  |  | 76.5278
| 1 | 5 | 73 | 74.817 | 77 | 77.9 | 73 | 76 | 75 | 76 | 75 | 76 | 78 | 75 | 77 | 75 | 73 | 74.109 | 76 | 75 | 74.8 | 75.507 | 81.5985 | 78.2 | 73.6468 | 78 |  |  | 147
| 1 | 6 | 78 | 75.7489 | 78 | 77.2 | 70 | 77 | 75 | 77 | 71 | 78 | 82 | 74 | 78 | 76 | 77 | 76.0949 | 77 | 79.1 | 75.6 | 74.03 | 80.535 | 78.7 | 71.7141 | 77
| Mean |  | 75.1667 | 74.0397 | 77.5 | 78.6 | 71 | 76.6667 | 74.8333 | 76.5 | 73.5 | 76.8333 | 79.1667 | 75.1667 | 77.1667 | 75.6667 | 75.6667 | 77.2454 | 76.5 | 78.1167 | 74.8667 | 77.9833 | 81.6643 | 78.4667 | 73.3529 | 76.8333
| Median |  | 75.5 | 73.5323 | 77.5 | 78.6 | 70 | 76.5 | 75 | 77 | 74 | 76 | 79.5 | 74.5 | 77 | 75.5 | 76 | 76.8768 | 76.5 | 78.55 | 74.8 | 76.0375 | 81.4343 | 78.5 | 73.3206 | 77
| Std Dev. |  | 2.3166 | 1.0145 | 0.54772 | 1.01 | 2.4495 | 1.7512 | 1.169 | 1.3784 | 2.0736 | 1.8348 | 2.0412 | 2.7142 | 1.3292 | 1.2111 | 1.9664 | 2.5157 | 0.54772 | 1.6092 | 0.75807 | 4.5586 | 1.0494 | 0.3266 | 2.0467 | 0.75277
| Rel.Std.Dev. |  | 3.082% | 1.37% | 0.7067% | 1.285% | 3.45% | 2.284% | 1.562% | 1.802% | 2.821% | 2.388% | 2.578% | 3.611% | 1.722% | 1.601% | 2.599% | 3.257% | 0.716% | 2.06% | 1.013% | 5.846% | 1.285% | 0.4162% | 2.79% | 0.9797%
| PDM3 |  | -1.779% | -3.251% | 1.27% | 2.708% | -7.223% | 0.1815% | -2.214% | -0.03632% | -3.956% | 0.3993% | 3.448% | -1.779% | 0.8348% | -1.125% | -1.125% | 0.9377% | -0.03632% | 2.076% | -2.171% | 1.902% | 6.712% | 2.534% | -4.149% | 0.3993%
| Z-Score (Absolute) |  | 0.67 | 1.21 | 0.44 | 0.96 | 2.66 | 0.04 | 0.83 | 0.04 | 1.47 | 0.12 | 1.23 | 0.67 | 0.28 | 0.44 | 0.44 | 0.32 | 0.04 | 0.73 | 0.82 | 0.67 | 2.42 | 0.9 | 1.54 | 0.12
| Analytical results for Zr in OREAS 471 (Certified Value 168 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 155 | 154.6491 | 143.1 | 171 | 149.7 | 145 | 173 | 165.1 | 117 | 187.3 | 173.5 | 167 | 172 | 172 | 174 | 328.7232 | 170 | 163.6 | 168.5 | 151.934 | 188.63 | 174.6 | 181.1393 | 191.1 |  |  | 1
| 1 | 2 | 156 | 155.0648 | 145.2 | 175 | 160.9 | 155 | 179 | 160.8 | 121 | 164.6 | 173 | 175 | 167.5 | 177 | 174 | 324.9268 | 170.9 | 169 | 166.8 | 159.863 | 186.59 | 175.9 | 180.7593 | 196.5 |  |  | #N/A
| 1 | 3 | 156 | 156.4119 | 140.4 | 166 | 158.2 | 156 | 175 | 170.4 | 120 | 170.5 | 168 | 166 | 161.5 | 179 | 172 | 333.5391 | 164.1 | 163.2 | 168.8 | 162.44 | 186.61 | 176.4 | 181.5671 | 192.1 |  |  | 16
| 1 | 4 | 167 | 155.1797 | 143.5 | 165 | 152.2 | 152 | 176 | 161 | 125 | 181.9 | 174.5 | 168.5 | 167 | 170.5 | 173.5 | 322.0195 | 172.7 | 167.1 | 169.2 | 155.819 | 187.81 | 175.2 | 179.0515 | 198.9 |  |  | 168.047
| 1 | 5 | 160 | 157.4031 | 145.9 | 169 | 166.3 | 149 | 172 | 169.1 | 124 | 157.8 | 170 | 168.5 | 168 | 173 | 168.5 | 323.16 | 166.2 | 165.5 | 169.3 | 157.991 | 184.72 | 175.5 | 183.7785 | 193.9 |  |  | 148
| 1 | 6 | 151 | 160.4865 | 143.4 | 165 | 156.5 | 154 | 177 | 166.5 | 113 | 168.4 | 177.5 | 168 | 166 | 173 | 171 | 325.9862 | 169.8 | 163.4 | 168.1 | 149.573 | 184.65 | 175.1 | 179.4233 | 199
| Mean |  | 157.5 | 156.5325 | 143.5833 | 168.5 | 157.3 | 151.8333 | 175.3333 | 165.4833 | 120 | 171.75 | 172.75 | 168.8333 | 167 | 174.0833 | 172.1667 | 326.3925 | 168.95 | 165.3 | 168.45 | 156.27 | 186.5017 | 175.45 | 180.9532 | 195.25
| Median |  | 156 | 155.7958 | 143.45 | 167.5 | 157.35 | 153 | 175.5 | 165.8 | 120.5 | 169.45 | 173.25 | 168.25 | 167.25 | 173 | 172.75 | 325.4565 | 169.9 | 164.55 | 168.65 | 156.905 | 186.6 | 175.35 | 180.9493 | 195.2
| Std Dev. |  | 5.4681 | 2.188 | 1.9177 | 3.9875 | 5.9863 | 4.1673 | 2.582 | 4.0117 | 4.4721 | 10.9849 | 3.3578 | 3.1728 | 3.3912 | 3.2314 | 2.1602 | 4.2046 | 3.1867 | 2.3614 | 0.9225 | 4.8538 | 1.604 | 0.63482 | 1.6964 | 3.4034
| Rel.Std.Dev. |  | 3.472% | 1.398% | 1.336% | 2.366% | 3.806% | 2.745% | 1.473% | 2.424% | 3.727% | 6.396% | 1.944% | 1.879% | 2.031% | 1.856% | 1.255% | 1.288% | 1.886% | 1.429% | 0.5476% | 3.106% | 0.86% | 0.3618% | 0.9375% | 1.743%
| PDM3 |  | -6.276% | -6.852% | -14.56% | 0.2695% | -6.395% | -9.648% | 4.336% | -1.526% | -28.59% | 2.204% | 2.799% | 0.4679% | -0.6231% | 3.592% | 2.451% | 94.23% | 0.5373% | -1.635% | 0.2398% | -7.008% | 10.98% | 4.405% | 7.68% | 16.19%
| Z-Score (Absolute) |  | 1.12 | 1.22 | 2.52 | 0.02 | 1.14 | 1.69 | 0.67 | 0.32 | 4.88 | 0.31 | 0.41 | 0.02 | 0.17 | 0.54 | 0.35 | 15.82 | 0.03 | 0.34 | 0.02 | 1.24 | 1.79 | 0.68 | 1.23 | 2.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C550: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C569: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C661: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C788: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C807: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C825: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C862: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C973: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C991: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1010: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1100: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.